## Abbreviations
Table 1. Abbreviations used for OREAS 600
Abbreviation | Explanation
Std.Dev. | one sigma standard deviation
Rel.Std.Dev. | one sigma relative standard deviation
PDM3 | percent deviation of lab mean from corrected mean of means
NR | not reported
< | ‘less than’
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
AD*AAS | acid digest (non-specific) with AAS finish
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
FA*AAS | fire assay with AAS finish
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
NB. Formatted area for below table ends at row 46 inclusive, allowing for 40 maximum potential laboratories.
Table 5. Participating Laboratory List used for OREAS 600
Lab Slot | Laboratory Name and Location
1.21 | Accurassay, Thunder Bay, Ontario, Canada
1.23 | Acme (BV), Santiago, Chile
2.01 | Actlabs, Ancaster, Ontario, Canada
1.2 | AH Knight, Spartanburg, SC, USA
1.06 | ALS, Johannesburg, South Africa
1.05 | ALS, Lima, Peru
1.03 | ALS, Reno, Nevada, USA
1.08 | ALS, Townsville, QLD, Australia
1.04 | ALS, Val-d'or, Quebec, Canada
1.07 | ALS, Vancouver, BC, Canada
1.24 | Bureau Veritas Geoanalytical, Adelaide, SA, Australia
1.27 | Bureau Veritas Geoanalytical, Kalgoorlie, WA, Australia
1.25 | Bureau Veritas Geoanalytical, Perth, WA, Australia
1.26 | Bureau Veritas Kalassay, Kalgoorlie, WA, Australia
1.11 | Inspectorate (BV), Lima, Peru
1.12 | Inspectorate (BV), Sparks, Nevada, USA
1.02 | Intertek Genalysis, Adelaide, SA, Australia
1.14 | Intertek Genalysis, Perth, WA, Australia
1.09 | Intertek Testing Services, Cupang, Muntinlupa, Philippines
1.15 | Intertek Testing Services, Shunyi, Beijing, China
1.22 | MINTEK Analytical Services, Randburg, South Africa
1.01 | PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
1.16 | SGS de Mexico, Durango, Mexico
1.18 | SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
1.19 | SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
1.17 | SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
1.13 | Shiva Analyticals Ltd, Bangalore North, Karnataka, India
1.1 | SRL, Perth, WA, Australia

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 600
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.20014 | 0.0056792 | 0.19817 | 0.20211 | 0.19653 | 0.20374
4-Acid Digestion
Ag, Silver (ppm) | 24.7513 | 1.0145 | 24.3092 | 25.1933 | 24.0655 | 25.437
Al, Aluminium (wt.%) | 6.7802 | 0.25215 | 6.6628 | 6.8976 | 6.541 | 7.0194
As, Arsenic (ppm) | 88.8006 | 7.2426 | 85.2176 | 92.3837 | 84.0429 | 93.5583
Be, Beryllium (ppm) | 1.6914 | 0.17634 | 1.6079 | 1.7748 | 1.5907 | 1.7921
Bi, Bismuth (ppm) | 6.3898 | 0.81184 | 6.0468 | 6.7327 | 6.0062 | 6.7733
Ca, Calcium (wt.%) | 1.8781 | 0.072883 | 1.8462 | 1.91 | 1.8375 | 1.9188
Cd, Cadmium (ppm) | 3.3731 | 0.31977 | 3.2379 | 3.5084 | 3.1703 | 3.576
Ce, Cerium (ppm) | 47.7034 | 1.6989 | 46.2665 | 49.1404 | 46.143 | 49.2639
Co, Cobalt (ppm) | 7.0554 | 0.84931 | 6.646 | 7.4647 | 6.7002 | 7.4106
Cr, Chromium (ppm) | 27.5108 | 4.2463 | 25.4725 | 29.5491 | 23.4372 | 31.5843
Cs, Cesium (ppm) | 9.4182 | 0.65076 | 8.9718 | 9.8645 | 9.0404 | 9.7959
Cu, Copper (ppm) | 481.9394 | 22.5856 | 471.9757 | 491.9032 | 469.9018 | 493.9771
Dy, Dysprosium (ppm) | 2.0534 | 0.11678 | 1.9323 | 2.1745 | 1.957 | 2.1498
Er, Erbium (ppm) | 0.96798 | 0.065152 | 0.88181 | 1.0542 | 0.92249 | 1.0135
Eu, Europium (ppm) | 1.0292 | 0.18564 | 0.82932 | 1.2292 | 0.98291 | 1.0756
Fe, Iron (wt.%) | 2.3757 | 0.11806 | 2.3215 | 2.4299 | 2.3196 | 2.4318
Ga, Gallium (ppm) | 18.7278 | 1.3122 | 18.0729 | 19.3828 | 17.805 | 19.6507
Gd, Gadolinium (ppm) | 3.1244 | 0.44897 | 2.6206 | 3.6281 | 2.9478 | 3.301
Hf, Hafnium (ppm) | 2.896 | 0.11818 | 2.8103 | 2.9818 | 2.7799 | 3.0121
Ho, Holmium (ppm) | 0.35548 | 0.025877 | 0.32232 | 0.38865 | 0.32374 | 0.38723
In, Indium (ppm) | 0.75032 | 0.070018 | 0.69879 | 0.80184 | 0.6842 | 0.81644
K, Potassium (wt.%) | 1.804 | 0.067358 | 1.7718 | 1.8361 | 1.7338 | 1.8741
La, Lanthanum (ppm) | 22.9505 | 2.9247 | 21.5913 | 24.3098 | 22.2281 | 23.673
Li, Lithium (ppm) | 18.6871 | 1.3051 | 17.9514 | 19.4228 | 18.0273 | 19.3469
Lu, Lutetium (ppm) | 0.12628 | 0.0093967 | 0.11516 | 0.1374 | IND | IND
Mg, Magnesium (wt.%) | 0.77164 | 0.034914 | 0.75597 | 0.78731 | 0.75645 | 0.78684
Mn, Manganese (wt.%) | 0.071033 | 0.0037504 | 0.069388 | 0.072678 | 0.069451 | 0.072615
Mo, Molybdenum (ppm) | 2.2033 | 0.4643 | 1.9942 | 2.4125 | 1.9662 | 2.4405
Na, Sodium (wt.%) | 0.58651 | 0.026781 | 0.57377 | 0.59926 | 0.57403 | 0.599
Nb, Niobium (ppm) | 7.9789 | 0.50967 | 7.6711 | 8.2867 | 7.6189 | 8.3389
Nd, Neodymium (ppm) | 19.4882 | 0.97019 | 18.468 | 20.5085 | 18.861 | 20.1155
Ni, Nickel (ppm) | 16.5011 | 2.7996 | 15.3149 | 17.6874 | 13.8646 | 19.1377
P, Phosphorus (wt.%) | 0.05962 | 0.0013662 | 0.059155 | 0.060085 | 0.057393 | 0.061847
Pb, Lead (ppm) | 193.085 | 13.6013 | 187.0753 | 199.0947 | 186.2252 | 199.9447
Pr, Praseodymium (ppm) | 5.4201 | 0.33951 | 4.9717 | 5.8684 | 5.2476 | 5.5925
Rb, Rubidium (ppm) | 75.2896 | 5.458 | 71.5173 | 79.062 | 73.1876 | 77.3917
S, Sulphur (wt.%) | 1.691 | 0.038663 | 1.6766 | 1.7054 | 1.6416 | 1.7404
Sb, Antimony (ppm) | 14.3374 | 1.8191 | 13.6223 | 15.0526 | 13.8095 | 14.8653
Sc, Scandium (ppm) | 5.9507 | 0.21925 | 5.8406 | 6.0607 | 5.5764 | 6.3249
Se, Selenium (ppm) | 6.974 | 1.3538 | 5.998 | 7.95 | 6.0572 | 7.8908
Sm, Samarium (ppm) | 3.5594 | 0.22541 | 3.3251 | 3.7936 | 3.3846 | 3.7341
Sn, Tin (ppm) | 2.117 | 0.19593 | 1.9545 | 2.2795 | 1.9707 | 2.2632
Sr, Strontium (ppm) | 186.3569 | 8.4308 | 182.6847 | 190.0291 | 181.7432 | 190.9706
Ta, Tantalum (ppm) | 0.63323 | 0.11456 | 0.51849 | 0.74796 | 0.57504 | 0.69141
Tb, Terbium (ppm) | 0.3967 | 0.052414 | 0.34655 | 0.44684 | 0.36829 | 0.4251
Te, Tellurium (ppm) | 7.3446 | 0.58536 | 6.9157 | 7.7735 | 6.9214 | 7.7678
Th, Thorium (ppm) | 9.3845 | 0.51374 | 8.9654 | 9.8036 | 8.9873 | 9.7817
Ti, Titanium (wt.%) | 0.24219 | 0.014424 | 0.23534 | 0.24904 | 0.23475 | 0.24963
Tl, Thallium (ppm) | 1.1129 | 0.058997 | 1.0838 | 1.142 | 1.0534 | 1.1724
U, Uranium (ppm) | 2.6866 | 0.15022 | 2.579 | 2.7942 | 2.5905 | 2.7827
V, Vanadium (ppm) | 45.4222 | 2.3148 | 44.4446 | 46.3999 | 43.7048 | 47.1397
W, Tungsten (ppm) | 4.1975 | 0.6777 | 3.7034 | 4.6916 | 3.8962 | 4.4988
Y, Yttrium (ppm) | 9.9267 | 0.48072 | 9.6398 | 10.2137 | 9.4864 | 10.367
Yb, Ytterbium (ppm) | 0.91257 | 0.064265 | 0.85807 | 0.96708 | 0.87056 | 0.95459
Zn, Zinc (ppm) | 614.5297 | 23.2257 | 603.8671 | 625.1922 | 590.911 | 638.1483
Zr, Zirconium (ppm) | 93.7526 | 2.9915 | 92.0395 | 95.4657 | 91.261 | 96.2443
Aqua Regia Digestion
Ag, Silver (ppm) | 24.3118 | 0.89812 | 23.8712 | 24.7524 | 23.714 | 24.9096
Al, Aluminium (wt.%) | 0.98431 | 0.14727 | 0.90755 | 1.0611 | 0.95654 | 1.0121
As, Arsenic (ppm) | 85.3495 | 7.0845 | 82.2479 | 88.4512 | 82.9361 | 87.763
Au, Gold (ppm) | 0.19246 | 0.010904 | 0.1871 | 0.19781 | 0.18866 | 0.19625
B, Boron (ppm) | < 10 | IND | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.70055 | 0.057837 | 0.66656 | 0.73455 | 0.6617 | 0.7394
Bi, Bismuth (ppm) | 6.0949 | 1.1356 | 5.6304 | 6.5594 | 5.7859 | 6.4038
Ca, Calcium (wt.%) | 1.7567 | 0.11908 | 1.7025 | 1.8109 | 1.7262 | 1.7872
Cd, Cadmium (ppm) | 3.4951 | 0.32676 | 3.3382 | 3.652 | 3.3246 | 3.6657
Ce, Cerium (ppm) | 34.804 | 2.8874 | 32.8652 | 36.7428 | 33.6903 | 35.9178
Co, Cobalt (ppm) | 6.6857 | 0.588 | 6.4489 | 6.9225 | 6.3914 | 6.98
Cr, Chromium (ppm) | 23.4036 | 3.0458 | 22.2525 | 24.5546 | 21.2148 | 25.5923
Cs, Cesium (ppm) | 2.6812 | 0.53222 | 2.1565 | 3.2059 | 2.5749 | 2.7874
Cu, Copper (ppm) | 488.0741 | 19.439 | 480.1163 | 496.0319 | 477.9133 | 498.235
Dy, Dysprosium (ppm) | 1.265 | 0.062251 | 1.2026 | 1.3273 | 1.2161 | 1.3138
Er, Erbium (ppm) | 0.53736 | 0.026693 | 0.5037 | 0.57102 | 0.51619 | 0.55852
Eu, Europium (ppm) | 0.58757 | 0.039656 | 0.53692 | 0.63822 | 0.56374 | 0.61141
Fe, Iron (wt.%) | 2.2196 | 0.11025 | 2.1704 | 2.2688 | 2.1788 | 2.2603
Ga, Gallium (ppm) | 4.1565 | 0.70335 | 3.6734 | 4.6396 | 3.9299 | 4.3831
Gd, Gadolinium (ppm) | 1.9691 | 0.14079 | 1.8125 | 2.1256 | 1.9132 | 2.025
Hf, Hafnium (ppm) | 0.21581 | 0.042259 | 0.16961 | 0.26201 | 0.18699 | 0.24463
Hg, Mercury (ppm) | < 1 | IND | IND | IND | IND | IND
Ho, Holmium (ppm) | 0.21894 | 0.0076263 | 0.21215 | 0.22573 | 0.21024 | 0.22764
In, Indium (ppm) | 0.7309 | 0.10335 | 0.68056 | 0.78124 | 0.69229 | 0.76952
K, Potassium (wt.%) | 0.20886 | 0.027844 | 0.19441 | 0.22331 | 0.20033 | 0.21739
La, Lanthanum (ppm) | 18.2764 | 1.4347 | 17.6035 | 18.9494 | 17.7334 | 18.8194
Li, Lithium (ppm) | 6.5375 | 1.0426 | 5.8945 | 7.1806 | 6.2391 | 6.8359
Lu, Lutetium (ppm) | 0.062557 | 0.0048219 | 0.057734 | 0.06738 | IND | IND
Mg, Magnesium (wt.%) | 0.3421 | 0.057331 | 0.31572 | 0.36848 | 0.33183 | 0.35237
Mn, Manganese (wt.%) | 0.068354 | 0.003482 | 0.066777 | 0.069931 | 0.067102 | 0.069606
Mo, Molybdenum (ppm) | 1.9153 | 0.31247 | 1.7803 | 2.0503 | 1.6886 | 2.142
Na, Sodium (wt.%) | 0.050627 | 0.0067625 | 0.047155 | 0.054099 | IND | IND
Nd, Neodymium (ppm) | 14.0837 | 1.0477 | 12.9183 | 15.249 | 13.6731 | 14.4943
Ni, Nickel (ppm) | 15.354 | 1.756 | 14.661 | 16.047 | 13.3958 | 17.3122
P, Phosphorus (wt.%) | 0.049827 | 0.0032192 | 0.048673 | 0.050981 | 0.048153 | 0.051502
Pb, Lead (ppm) | 156.6304 | 4.9534 | 154.9876 | 158.2732 | 152.0901 | 161.1707
Pr, Praseodymium (ppm) | 4.0351 | 0.14391 | 3.8527 | 4.2175 | 3.9048 | 4.1654
S, Sulphur (wt.%) | 1.6587 | 0.061728 | 1.6272 | 1.6902 | 1.6303 | 1.687
Sb, Antimony (ppm) | < 15 | IND | IND | IND | IND | IND
Sc, Scandium (ppm) | 2.0607 | 0.23638 | 1.9386 | 2.1828 | 1.9484 | 2.1731
Se, Selenium (ppm) | 6.7749 | 0.82886 | 6.2804 | 7.2694 | 6.3278 | 7.2221
Sm, Samarium (ppm) | 2.4886 | 0.18395 | 2.2862 | 2.691 | 2.3903 | 2.5869
Sn, Tin (ppm) | 1.179 | 0.12241 | 1.0621 | 1.296 | 1.1107 | 1.2474
Sr, Strontium (ppm) | 38.3771 | 4.0767 | 36.4609 | 40.2934 | 37.342 | 39.4122
Tb, Terbium (ppm) | 0.25423 | 0.021348 | 0.23618 | 0.27227 | 0.24083 | 0.26762
Te, Tellurium (ppm) | 6.7983 | 0.74029 | 6.2424 | 7.3541 | 6.5459 | 7.0506
Th, Thorium (ppm) | 4.9082 | 0.30591 | 4.6738 | 5.1427 | 4.7152 | 5.1013
Ti, Titanium (wt.%) | < 0.01 | IND | IND | IND | IND | IND
Tl, Thallium (ppm) | 0.562 | 0.051933 | 0.52362 | 0.60038 | 0.54365 | 0.58035
U, Uranium (ppm) | 1.0248 | 0.097721 | 0.95256 | 1.097 | 0.98088 | 1.0687
V, Vanadium (ppm) | 12.9328 | 2.1467 | 11.8677 | 13.9979 | 12.1407 | 13.7249
Y, Yttrium (ppm) | 5.9785 | 0.42626 | 5.7017 | 6.2552 | 5.7568 | 6.2001
Yb, Ytterbium (ppm) | 0.42656 | 0.024786 | 0.41677 | 0.43634 | 0.40853 | 0.44459
Zn, Zinc (ppm) | 597.7687 | 35.2646 | 582.3753 | 613.162 | 584.1794 | 611.3579
Zr, Zirconium (ppm) | 10.2144 | 1.9684 | 8.8643 | 11.5645 | 9.9536 | 10.4752
Infrared Combustion
S, Sulphur (wt.%) | 1.6686 | 0.040237 | 1.6458 | 1.6914 | 1.6414 | 1.6958

## Indicative Values
Table 3. Indicative Values for OREAS 600
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Ag | ppm | 13.9667 | Pd | ppb | < 5 | Pt | ppb | < 5
4-Acid Digestion
B | ppm | < 20 | Ge | ppm | 0.78035 | Re | ppb | < 2
Ba | ppm | 1844.6411 | Hg | ppm | < 1 | Tm | ppm | 0.12242
Aqua Regia Digestion
Ba | ppm | 1288.75 | Pt | ppb | < 5 | Si | wt.% | 0.036667
Ge | ppm | < 0.1 | Rb | ppm | 10.7639 | Ta | ppm | 0.12024
Nb | ppm | 0.17942 | Re | ppb | < 1 | Tm | ppm | 0.064324
Pd | ppb | < 10 | Ru | ppb | < 2 | W | ppm | 0.45695
Infrared Combustion
C | wt.% | 0.48818

## Indicative Values (2)
Table 3. Indicative Values for OREAS 600
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Laser Ablation ICP-MS
Ag | ppm | 20.85 | Ho | ppm | 0.355 | Sn | ppm | 3
As | ppm | 84.9 | In | ppm | 0.725 | Sr | ppm | 189.5
Ba | ppm | 1855 | La | ppm | 26.95 | Ta | ppm | 0.665
Be | ppm | 0.8 | Lu | ppm | 0.135 | Tb | ppm | 0.345
Bi | ppm | 6.23 | Mn | wt.% | 0.0823 | Te | ppm | 7.2
Cd | ppm | 2.95 | Mo | ppm | 2.4 | Th | ppm | 9.92
Ce | ppm | 48.05 | Nb | ppm | 8.265 | Ti | wt.% | 0.249
Co | ppm | 7.1 | Nd | ppm | 20 | Tl | ppm | 1.1
Cr | ppm | 46 | Ni | ppm | 17 | Tm | ppm | 0.15
Cs | ppm | 9.335 | Pb | ppm | 190 | U | ppm | 2.755
Cu | ppm | 461 | Pr | ppm | 5.585 | V | ppm | 47.05
Dy | ppm | 2.04 | Rb | ppm | 72.7 | W | ppm | 3.675
Er | ppm | 1.025 | Re | ppm | 0.0075 | Y | ppm | 10.75
Eu | ppm | 0.85 | Sb | ppm | 13.55 | Yb | ppm | 0.965
Ga | ppm | 18.35 | Sc | ppm | 6 | Zn | ppm | 580
Gd | ppm | 2.91 | Se | ppm | < 5 | Zr | ppm | 134
Hf | ppm | 3.775 | Sm | ppm | 3.705
Borate Fusion XRF
Al2O3 | wt.% | 13.39 | Fe2O3 | wt.% | 3.435 | Pb | ppm | 190
As | ppm | 100 | K2O | wt.% | 2.235 | SiO2 | wt.% | 68.045
Ba | ppm | 1865 | MgO | wt.% | 1.355 | Sn | ppm | 12.5
CaO | wt.% | 2.695 | MnO | wt.% | 0.105 | SO3 | wt.% | 4.1085
Co | ppm | 10 | Na2O | wt.% | 0.835 | TiO2 | wt.% | 0.417
Cr | ppm | 40 | Ni | ppm | < 10 | U | ppm | < 10
Cu | ppm | 475 | P2O5 | wt.% | 0.143 | Zn | ppm | 585
Thermogravimetry
LOI1000 | wt.% | 4.52

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 600
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.20014 | 0.0056792 | 0.18878 | 0.2115 | 0.1831 | 0.21718 | 2.838% | 5.675% | 8.513% | 0.19013 | 0.21015
4-Acid Digestion
Ag, ppm | 24.7513 | 1.0145 | 22.7224 | 26.7802 | 21.7079 | 27.7946 | 4.099% | 8.197% | 12.3% | 23.5137 | 25.9888
Al, wt.% | 6.7802 | 0.25215 | 6.2759 | 7.2845 | 6.0237 | 7.5366 | 3.719% | 7.438% | 11.16% | 6.4412 | 7.1192
As, ppm | 88.8006 | 7.2426 | 74.3155 | 103.2858 | 67.0729 | 110.5284 | 8.156% | 16.31% | 24.47% | 84.3606 | 93.2407
Be, ppm | 1.6914 | 0.17634 | 1.3387 | 2.0441 | 1.1624 | 2.2204 | 10.43% | 20.85% | 31.28% | 1.6068 | 1.7759
Bi, ppm | 6.3898 | 0.81184 | 4.7661 | 8.0134 | 3.9542 | 8.8253 | 12.71% | 25.41% | 38.12% | 6.0703 | 6.7092
Ca, wt.% | 1.8781 | 0.072883 | 1.7323 | 2.0239 | 1.6595 | 2.0968 | 3.881% | 7.761% | 11.64% | 1.7842 | 1.972
Cd, ppm | 3.3731 | 0.31977 | 2.7336 | 4.0127 | 2.4138 | 4.3325 | 9.48% | 18.96% | 28.44% | 3.2045 | 3.5418
Ce, ppm | 47.7034 | 1.6989 | 44.3055 | 51.1013 | 42.6066 | 52.8003 | 3.561% | 7.123% | 10.68% | 45.3183 | 50.0886
Co, ppm | 7.0554 | 0.84931 | 5.3568 | 8.754 | 4.5075 | 9.6033 | 12.04% | 24.08% | 36.11% | 6.7026 | 7.4081
Cr, ppm | 27.5108 | 4.2463 | 19.0182 | 36.0033 | 14.772 | 40.2496 | 15.43% | 30.87% | 46.3% | 26.1352 | 28.8863
Cs, ppm | 9.4182 | 0.65076 | 8.1166 | 10.7197 | 7.4659 | 11.3704 | 6.91% | 13.82% | 20.73% | 8.9472 | 9.8891
Cu, ppm | 481.9394 | 22.5856 | 436.7683 | 527.1106 | 414.1827 | 549.6962 | 4.686% | 9.373% | 14.06% | 457.8425 | 506.0364
Dy, ppm | 2.0534 | 0.11678 | 1.8198 | 2.2869 | 1.703 | 2.4037 | 5.687% | 11.37% | 17.06% | 1.9507 | 2.156
Er, ppm | 0.96798 | 0.065152 | 0.83768 | 1.0983 | 0.77253 | 1.1634 | 6.731% | 13.46% | 20.19% | 0.91958 | 1.0164
Eu, ppm | 1.0292 | 0.18564 | 0.65797 | 1.4005 | 0.47233 | 1.5862 | 18.04% | 36.07% | 54.11% | 0.97778 | 1.0807
Fe, wt.% | 2.3757 | 0.11806 | 2.1396 | 2.6118 | 2.0215 | 2.7299 | 4.969% | 9.939% | 14.91% | 2.2569 | 2.4945
Ga, ppm | 18.7278 | 1.3122 | 16.1035 | 21.3522 | 14.7913 | 22.6644 | 7.007% | 14.01% | 21.02% | 17.7914 | 19.6642
Gd, ppm | 3.1244 | 0.44897 | 2.2264 | 4.0223 | 1.7775 | 4.4713 | 14.37% | 28.74% | 43.11% | 2.9682 | 3.2806
Hf, ppm | 2.896 | 0.11818 | 2.6597 | 3.1324 | 2.5415 | 3.2506 | 4.081% | 8.162% | 12.24% | 2.7512 | 3.0408
Ho, ppm | 0.35548 | 0.025877 | 0.30373 | 0.40724 | 0.27785 | 0.43311 | 7.28% | 14.56% | 21.84% | 0.33771 | 0.37326
In, ppm | 0.75032 | 0.070018 | 0.61028 | 0.89035 | 0.54026 | 0.96037 | 9.332% | 18.66% | 28% | 0.7128 | 0.78783
K, wt.% | 1.804 | 0.067358 | 1.6693 | 1.9387 | 1.6019 | 2.0061 | 3.734% | 7.468% | 11.2% | 1.7138 | 1.8942
La, ppm | 22.9505 | 2.9247 | 17.1012 | 28.7999 | 14.1765 | 31.7246 | 12.74% | 25.49% | 38.23% | 21.803 | 24.0981
Li, ppm | 18.6871 | 1.3051 | 16.0769 | 21.2973 | 14.7718 | 22.6024 | 6.984% | 13.97% | 20.95% | 17.7528 | 19.6215
Lu, ppm | 0.12628 | 0.0093967 | 0.10749 | 0.14508 | 0.098092 | 0.15447 | 7.441% | 14.88% | 22.32% | 0.11997 | 0.1326
Mg, wt.% | 0.77164 | 0.034914 | 0.70182 | 0.84147 | 0.6669 | 0.87638 | 4.525% | 9.049% | 13.57% | 0.73306 | 0.81023
Mn, wt.% | 0.071033 | 0.0037504 | 0.063532 | 0.078534 | 0.059782 | 0.082284 | 5.28% | 10.56% | 15.84% | 0.067481 | 0.074585
Mo, ppm | 2.2033 | 0.4643 | 1.2747 | 3.1319 | 0.81041 | 3.5962 | 21.07% | 42.15% | 63.22% | 2.0932 | 2.3135
Na, wt.% | 0.58651 | 0.026781 | 0.53295 | 0.64007 | 0.50617 | 0.66685 | 4.566% | 9.132% | 13.7% | 0.55719 | 0.61584
Nb, ppm | 7.9789 | 0.50967 | 6.9596 | 8.9982 | 6.4499 | 9.5079 | 6.388% | 12.78% | 19.16% | 7.58 | 8.3778
Nd, ppm | 19.4882 | 0.97019 | 17.5479 | 21.4286 | 16.5777 | 22.3988 | 4.978% | 9.957% | 14.93% | 18.5138 | 20.4627
Ni, ppm | 16.5011 | 2.7996 | 10.9019 | 22.1004 | 8.1022 | 24.9001 | 16.97% | 33.93% | 50.9% | 15.6761 | 17.3262
P, wt.% | 0.05962 | 0.0013662 | 0.056888 | 0.062352 | 0.055521 | 0.063719 | 2.292% | 4.583% | 6.875% | 0.056639 | 0.062601
Pb, ppm | 193.085 | 13.6013 | 165.8824 | 220.2876 | 152.2811 | 233.8889 | 7.044% | 14.09% | 21.13% | 183.4307 | 202.7392
Pr, ppm | 5.4201 | 0.33951 | 4.741 | 6.0991 | 4.4015 | 6.4386 | 6.264% | 12.53% | 18.79% | 5.1491 | 5.6911
Rb, ppm | 75.2896 | 5.458 | 64.3736 | 86.2057 | 58.9156 | 91.6637 | 7.249% | 14.5% | 21.75% | 71.5252 | 79.0541
S, wt.% | 1.691 | 0.038663 | 1.6137 | 1.7683 | 1.575 | 1.807 | 2.286% | 4.573% | 6.859% | 1.6065 | 1.7756
Sb, ppm | 14.3374 | 1.8191 | 10.6993 | 17.9756 | 8.8802 | 19.7946 | 12.69% | 25.38% | 38.06% | 13.6206 | 15.0543
Sc, ppm | 5.9507 | 0.21925 | 5.5122 | 6.3891 | 5.2929 | 6.6084 | 3.684% | 7.369% | 11.05% | 5.6531 | 6.2482
Se, ppm | 6.974 | 1.3538 | 4.2663 | 9.6817 | 2.9125 | 11.0355 | 19.41% | 38.83% | 58.24% | 6.6253 | 7.3227
Sm, ppm | 3.5594 | 0.22541 | 3.1085 | 4.0102 | 2.8831 | 4.2356 | 6.333% | 12.67% | 19% | 3.3814 | 3.7373
Sn, ppm | 2.117 | 0.19593 | 1.7251 | 2.5088 | 1.5292 | 2.7048 | 9.255% | 18.51% | 27.77% | 2.0111 | 2.2228
Sr, ppm | 186.3569 | 8.4308 | 169.4953 | 203.2185 | 161.0645 | 211.6494 | 4.524% | 9.048% | 13.57% | 177.0391 | 195.6748
Ta, ppm | 0.63323 | 0.11456 | 0.40411 | 0.86234 | 0.28956 | 0.97689 | 18.09% | 36.18% | 54.27% | 0.60156 | 0.66489
Tb, ppm | 0.3967 | 0.052414 | 0.29187 | 0.50152 | 0.23945 | 0.55394 | 13.21% | 26.43% | 39.64% | 0.37686 | 0.41653
Te, ppm | 7.3446 | 0.58536 | 6.1739 | 8.5153 | 5.5885 | 9.1007 | 7.97% | 15.94% | 23.91% | 6.9774 | 7.7118
Th, ppm | 9.3845 | 0.51374 | 8.357 | 10.412 | 7.8433 | 10.9257 | 5.474% | 10.95% | 16.42% | 8.9153 | 9.8537
Ti, wt.% | 0.24219 | 0.014424 | 0.21334 | 0.27104 | 0.19892 | 0.28546 | 5.956% | 11.91% | 17.87% | 0.23008 | 0.2543
Tl, ppm | 1.1129 | 0.058997 | 0.99492 | 1.2309 | 0.93592 | 1.2899 | 5.301% | 10.6% | 15.9% | 1.0573 | 1.1686
U, ppm | 2.6866 | 0.15022 | 2.3862 | 2.9871 | 2.2359 | 3.1373 | 5.592% | 11.18% | 16.77% | 2.5523 | 2.8209
V, ppm | 45.4222 | 2.3148 | 40.7926 | 50.0519 | 38.4778 | 52.3667 | 5.096% | 10.19% | 15.29% | 43.1511 | 47.6934
W, ppm | 4.1975 | 0.6777 | 2.8421 | 5.5529 | 2.1644 | 6.2306 | 16.15% | 32.29% | 48.44% | 3.9876 | 4.4074
Y, ppm | 9.9267 | 0.48072 | 8.9653 | 10.8881 | 8.4846 | 11.3689 | 4.843% | 9.685% | 14.53% | 9.4304 | 10.423
Yb, ppm | 0.91257 | 0.064265 | 0.78404 | 1.0411 | 0.71978 | 1.1054 | 7.042% | 14.08% | 21.13% | 0.86694 | 0.9582
Zn, ppm | 614.5297 | 23.2257 | 568.0782 | 660.9811 | 544.8524 | 684.2069 | 3.779% | 7.559% | 11.34% | 583.8032 | 645.2561
Zr, ppm | 93.7526 | 2.9915 | 87.7697 | 99.7356 | 84.7782 | 102.727 | 3.191% | 6.382% | 9.572% | 89.065 | 98.4403
Aqua Regia Digestion
Ag, ppm | 24.3118 | 0.89812 | 22.5156 | 26.108 | 21.6175 | 27.0062 | 3.694% | 7.388% | 11.08% | 23.0962 | 25.5274
Al, wt.% | 0.98431 | 0.14727 | 0.68977 | 1.2789 | 0.5425 | 1.4261 | 14.96% | 29.92% | 44.89% | 0.9351 | 1.0335
As, ppm | 85.3495 | 7.0845 | 71.1805 | 99.5186 | 64.096 | 106.6031 | 8.301% | 16.6% | 24.9% | 81.0821 | 89.617
Au, ppm | 0.19246 | 0.010904 | 0.17065 | 0.21427 | 0.15975 | 0.22517 | 5.666% | 11.33% | 17% | 0.18284 | 0.20208
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.70055 | 0.057837 | 0.58488 | 0.81623 | 0.52704 | 0.87406 | 8.256% | 16.51% | 24.77% | 0.66552 | 0.73558
Bi, ppm | 6.0949 | 1.1356 | 3.8236 | 8.3661 | 2.688 | 9.5017 | 18.63% | 37.26% | 55.9% | 5.7901 | 6.3996
Ca, wt.% | 1.7567 | 0.11908 | 1.5185 | 1.9949 | 1.3995 | 2.114 | 6.779% | 13.56% | 20.34% | 1.6689 | 1.8445
Cd, ppm | 3.4951 | 0.32676 | 2.8416 | 4.1487 | 2.5148 | 4.4754 | 9.349% | 18.7% | 28.05% | 3.3204 | 3.6699
Ce, ppm | 34.804 | 2.8874 | 29.0293 | 40.5787 | 26.142 | 43.4661 | 8.296% | 16.59% | 24.89% | 33.0638 | 36.5442
Co, ppm | 6.6857 | 0.588 | 5.5097 | 7.8617 | 4.9217 | 8.4497 | 8.795% | 17.59% | 26.38% | 6.3514 | 7.02
Cr, ppm | 23.4036 | 3.0458 | 17.3121 | 29.4951 | 14.2663 | 32.5408 | 13.01% | 26.03% | 39.04% | 22.2334 | 24.5737
Cs, ppm | 2.6812 | 0.53222 | 1.6167 | 3.7456 | 1.0845 | 4.2778 | 19.85% | 39.7% | 59.55% | 2.5471 | 2.8152
Cu, ppm | 488.0741 | 19.439 | 449.1962 | 526.952 | 429.7572 | 546.391 | 3.983% | 7.966% | 11.95% | 463.6704 | 512.4778
Dy, ppm | 1.265 | 0.062251 | 1.1405 | 1.3895 | 1.0782 | 1.4517 | 4.921% | 9.842% | 14.76% | 1.2017 | 1.3282
Er, ppm | 0.53736 | 0.026693 | 0.48397 | 0.59075 | 0.45728 | 0.61744 | 4.968% | 9.935% | 14.9% | 0.51049 | 0.56423
Eu, ppm | 0.58757 | 0.039656 | 0.50826 | 0.66689 | 0.46861 | 0.70654 | 6.749% | 13.5% | 20.25% | 0.5582 | 0.61695
Fe, wt.% | 2.2196 | 0.11025 | 1.9991 | 2.4401 | 1.8888 | 2.5503 | 4.967% | 9.934% | 14.9% | 2.1086 | 2.3306
Ga, ppm | 4.1565 | 0.70335 | 2.7498 | 5.5632 | 2.0465 | 6.2666 | 16.92% | 33.84% | 50.76% | 3.9487 | 4.3644
Gd, ppm | 1.9691 | 0.14079 | 1.6875 | 2.2506 | 1.5467 | 2.3914 | 7.15% | 14.3% | 21.45% | 1.8706 | 2.0675
Hf, ppm | 0.21581 | 0.042259 | 0.13129 | 0.30033 | 0.089029 | 0.34259 | 19.58% | 39.16% | 58.75% | 0.20502 | 0.2266
Hg, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ho, ppm | 0.21894 | 0.0076263 | 0.20369 | 0.23419 | 0.19606 | 0.24182 | 3.483% | 6.967% | 10.45% | 0.20799 | 0.22989
In, ppm | 0.7309 | 0.10335 | 0.52421 | 0.9376 | 0.42086 | 1.0409 | 14.14% | 28.28% | 42.42% | 0.69436 | 0.76745
K, wt.% | 0.20886 | 0.027844 | 0.15317 | 0.26455 | 0.12533 | 0.29239 | 13.33% | 26.66% | 39.99% | 0.19842 | 0.2193
La, ppm | 18.2764 | 1.4347 | 15.4071 | 21.1458 | 13.9724 | 22.5805 | 7.85% | 15.7% | 23.55% | 17.3626 | 19.1903
Li, ppm | 6.5375 | 1.0426 | 4.4523 | 8.6228 | 3.4096 | 9.6654 | 15.95% | 31.9% | 47.85% | 6.2107 | 6.8644
Lu, ppm | 0.062557 | 0.0048219 | 0.052913 | 0.072201 | 0.048091 | 0.077023 | 7.708% | 15.42% | 23.12% | 0.059429 | 0.065685
Mg, wt.% | 0.3421 | 0.057331 | 0.22744 | 0.45676 | 0.17011 | 0.51409 | 16.76% | 33.52% | 50.28% | 0.32499 | 0.3592
Mn, wt.% | 0.068354 | 0.003482 | 0.06139 | 0.075318 | 0.057908 | 0.0788 | 5.094% | 10.19% | 15.28% | 0.064936 | 0.071772
Mo, ppm | 1.9153 | 0.31247 | 1.2903 | 2.5402 | 0.97787 | 2.8527 | 16.31% | 32.63% | 48.94% | 1.8195 | 2.011
Na, wt.% | 0.050627 | 0.0067625 | 0.037102 | 0.064152 | 0.030339 | 0.070914 | 13.36% | 26.72% | 40.07% | 0.048095 | 0.053158
Nd, ppm | 14.0837 | 1.0477 | 11.9883 | 16.1791 | 10.9406 | 17.2268 | 7.439% | 14.88% | 22.32% | 13.3795 | 14.7879
Ni, ppm | 15.354 | 1.756 | 11.842 | 18.8659 | 10.086 | 20.6219 | 11.44% | 22.87% | 34.31% | 14.5863 | 16.1217
P, wt.% | 0.049827 | 0.0032192 | 0.043389 | 0.056266 | 0.04017 | 0.059485 | 6.461% | 12.92% | 19.38% | 0.047336 | 0.052319
Pb, ppm | 156.6304 | 4.9534 | 146.7235 | 166.5372 | 141.7701 | 171.4907 | 3.162% | 6.325% | 9.487% | 148.7989 | 164.4619
Pr, ppm | 4.0351 | 0.14391 | 3.7473 | 4.3229 | 3.6034 | 4.4668 | 3.566% | 7.133% | 10.7% | 3.8334 | 4.2369
S, wt.% | 1.6587 | 0.061728 | 1.5352 | 1.7821 | 1.4735 | 1.8439 | 3.722% | 7.443% | 11.16% | 1.5757 | 1.7416
Sb, ppm | < 15 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sc, ppm | 2.0607 | 0.23638 | 1.588 | 2.5335 | 1.3516 | 2.7699 | 11.47% | 22.94% | 34.41% | 1.9577 | 2.1638
Se, ppm | 6.7749 | 0.82886 | 5.1172 | 8.4326 | 4.2883 | 9.2615 | 12.23% | 24.47% | 36.7% | 6.4362 | 7.1137
Sm, ppm | 2.4886 | 0.18395 | 2.1207 | 2.8565 | 1.9367 | 3.0405 | 7.392% | 14.78% | 22.18% | 2.3642 | 2.613
Sn, ppm | 1.179 | 0.12241 | 0.93422 | 1.4238 | 0.81181 | 1.5463 | 10.38% | 20.76% | 31.15% | 1.1201 | 1.238
Sr, ppm | 38.3771 | 4.0767 | 30.2236 | 46.5306 | 26.1469 | 50.6074 | 10.62% | 21.25% | 31.87% | 36.4583 | 40.296
Tb, ppm | 0.25423 | 0.021348 | 0.21153 | 0.29692 | 0.19018 | 0.31827 | 8.397% | 16.79% | 25.19% | 0.24151 | 0.26694
Te, ppm | 6.7983 | 0.74029 | 5.3177 | 8.2788 | 4.5774 | 9.0191 | 10.89% | 21.78% | 32.67% | 6.4583 | 7.1382
Th, ppm | 4.9082 | 0.30591 | 4.2964 | 5.5201 | 3.9905 | 5.826 | 6.233% | 12.47% | 18.7% | 4.6628 | 5.1537
Ti, wt.% | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tl, ppm | 0.562 | 0.051933 | 0.45813 | 0.66587 | 0.4062 | 0.7178 | 9.241% | 18.48% | 27.72% | 0.5339 | 0.5901
U, ppm | 1.0248 | 0.097721 | 0.82933 | 1.2202 | 0.73161 | 1.3179 | 9.536% | 19.07% | 28.61% | 0.97354 | 1.076
V, ppm | 12.9328 | 2.1467 | 8.6394 | 17.2262 | 6.4927 | 19.3729 | 16.6% | 33.2% | 49.8% | 12.2862 | 13.5794
Y, ppm | 5.9785 | 0.42626 | 5.1259 | 6.831 | 4.6997 | 7.2572 | 7.13% | 14.26% | 21.39% | 5.6795 | 6.2774
Yb, ppm | 0.42656 | 0.024786 | 0.37699 | 0.47613 | 0.3522 | 0.50091 | 5.811% | 11.62% | 17.43% | 0.40523 | 0.44788
Zn, ppm | 597.7687 | 35.2646 | 527.2395 | 668.2978 | 491.975 | 703.5623 | 5.899% | 11.8% | 17.7% | 567.8802 | 627.6571
Zr, ppm | 10.2144 | 1.9684 | 6.2777 | 14.1512 | 4.3093 | 16.1195 | 19.27% | 38.54% | 57.81% | 9.7037 | 10.7251
Infrared Combustion
S, wt.% | 1.6686 | 0.040237 | 1.5881 | 1.7491 | 1.5479 | 1.7893 | 2.411% | 4.823% | 7.234% | 1.5852 | 1.752

## Fire Assay
| Analytical results for Ag in OREAS 600 (Indicative Value 14 ppm) | [29 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 01 | 11 | 12 | 16 | 17 | 19 | 23 | [21 empty] | ppm
|  |  | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | [21 empty] | 1
|  |  | 40g | 30g | 30g | 30g | 30g | 30g | 30g | [21 empty] | 1
| 1 | 1 | 8.8 | N.A. | N.A. | N.A. | - | --- | N.A. | [21 empty] | 1
| 1 | 2 | 5.8 | N.A. | < 50 | N.A. | - | --- | N.A. | [21 empty] | 1
| 1 | 3 | 12.8 | N.A. | N.A. | N.A. | - | --- | N.A. | [21 empty] | 16
| 1 | 4 | 23.3 | N.A. | N.A. | N.A. | - | --- | N.A. | [21 empty] | 13.9667
| 1 | 5 | 18.3 | N.A. | N.A. | N.A. | - | --- | N.A. | [21 empty] | 7
| 1 | 6 | 14.8 | N.A. | N.A. | N.A. | - | --- | N.A.
| Mean |  | 13.9667
| Median |  | 13.8
| Std Dev. |  | 6.3456
| Rel.Std.Dev. |  | 45.43%
| PDM3 |  | -2.442e-13%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Au in OREAS 600 (Certified Value 0.2 ppm) | [29 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | 3
|  |  | 1.0g | 40g | 25g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 30g | 20g | 40g | 40g | 20g | 3
| 1 | 1 | 0.193 | 0.197 | 0.2 | 0.203 | 0.191 | 0.192 | 0.199 | 0.192 | 0.189 | 0.20908 | 0.2 | 0.197 | 0.19 | 0.211 | 0.19 | 0.2 | 0.2 | 0.2 | 0.2 | 0.216 | 0.2 | 0.195 |  | 0.203 | 0.2 | 0.198 | 0.21 | 0.2 | 1
| 1 | 2 | 0.199 | 0.201 | 0.2 | 0.206 | 0.191 | 0.191 | 0.211 | 0.188 | 0.193 | 0.2085 | 0.2 | 0.191 | 0.199 | 0.222 | 0.19 | 0.2 | 0.196 | 0.19 | 0.2 | 0.205 | 0.2 | 0.203 |  | 0.201 | 0.2 | 0.198 | 0.2 | 0.2 | 0
| 1 | 3 | 0.207 | 0.207 | 0.2 | 0.202 | 0.201 | 0.188 | 0.206 | 0.198 | 0.187 | 0.20867 | 0.2 | 0.198 | 0.201 | 0.215 | 0.19 | 0.19 | 0.202 | 0.22 | 0.2 | 0.203 | 0.2 | 0.188 |  | 0.204 | 0.2 | 0.204 | 0.2 | 0.18 | 16
| 1 | 4 | 0.193 | 0.211 | 0.2 | 0.203 | 0.192 | 0.197 | 0.2 | 0.195 | 0.185 | 0.2085 | 0.2 | 0.198 | 0.196 | 0.211 | 0.19 | 0.21 | 0.194 | 0.22 | 0.2 | 0.202 | 0.2 | 0.193 | 0.22 | 0.201 | 0.2 | 0.2 | 0.21 | 0.19 | 0.20014
| 1 | 5 | 0.209 | 0.203 | 0.19 | 0.197 | 0.197 | 0.201 | 0.203 | 0.193 | 0.18 | 0.20783 | 0.2 | 0.198 | 0.197 | 0.212 | 0.18 | 0.21 | 0.199 | 0.22 | 0.2 | 0.201 | 0.2 | 0.19 | 0.21 | 0.202 | 0.18 | 0.202 | 0.2 | 0.195 | 7
| 1 | 6 | 0.195 | 0.201 | 0.19 | 0.219 | 0.196 | 0.207 | 0.204 | 0.191 | 0.182 | 0.20817 | 0.2 | 0.203 | 0.198 | 0.21 | 0.18 | 0.2 | 0.196 | 0.24 | 0.2 | 0.206 | 0.2 | 0.195 | 0.22 | 0.22 | 0.2 | 0.2 | 0.2 | 0.19
|  | 7 | 00: 0.203
|  | 8 | 00: 0.195
|  | 9 | 00: 0.191
|  | 10 | 00: 0.199
|  | 11 | 00: 0.201
|  | 12 | 00: 0.209
|  | 13 | 00: 0.19
|  | 14 | 00: 0.202
|  | 15 | 00: 0.201
|  | 16 | 00: 0.186
|  | 17 | 00: 0.198
|  | 18 | 00: 0.2
|  | 19 | 00: 0.198
|  | 20 | 00: 0.197
| Mean |  | 0.1983 | 0.20333 | 0.19667 | 0.205 | 0.19467 | 0.196 | 0.20383 | 0.19283 | 0.186 | 0.20846 | 0.2 | 0.1975 | 0.19683 | 0.2135 | 0.18667 | 0.20167 | 0.19783 | 0.215 | 0.2 | 0.2055 | 0.2 | 0.194 | 0.21667 | 0.20517 | 0.19667 | 0.20033 | 0.20333 | 0.1925
| Median |  | 0.1985 | 0.202 | 0.2 | 0.203 | 0.194 | 0.1945 | 0.2035 | 0.1925 | 0.186 | 0.2085 | 0.2 | 0.198 | 0.1975 | 0.2115 | 0.19 | 0.2 | 0.1975 | 0.22 | 0.2 | 0.204 | 0.2 | 0.194 | 0.22 | 0.2025 | 0.2 | 0.2 | 0.2 | 0.1925
| Std Dev. |  | 0.0061224 | 0.0049666 | 0.005164 | 0.0074565 | 0.0040332 | 0.0070993 | 0.0043551 | 0.0034303 | 0.0047329 | 0.00042736 | 3.0405e-17 | 0.0038341 | 0.0037639 | 0.0045056 | 0.005164 | 0.0075277 | 0.0029944 | 0.017607 | 3.0405e-17 | 0.0054681 | 3.0405e-17 | 0.0052154 | 0.0057735 | 0.0073598 | 0.008165 | 0.0023381 | 0.005164 | 0.0075829
| Rel.Std.Dev. |  | 3.087% | 2.443% | 2.626% | 3.637% | 2.072% | 3.622% | 2.137% | 1.779% | 2.545% | 0.205% | 1.52e-14% | 1.941% | 1.912% | 2.11% | 2.766% | 3.733% | 1.514% | 8.189% | 1.52e-14% | 2.661% | 1.52e-14% | 2.688% | 2.665% | 3.587% | 4.152% | 1.167% | 2.54% | 3.939%
| PDM3 |  | -0.9187% | 1.596% | -1.735% | 2.429% | -2.734% | -2.068% | 1.846% | -3.65% | -7.064% | 4.157% | -0.06934% | -1.318% | -1.652% | 6.676% | -6.731% | 0.7634% | -1.152% | 7.425% | -0.06934% | 2.679% | -0.06934% | -3.067% | 8.258% | 2.512% | -1.735% | 0.09722% | 1.596% | -3.817%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67 | 0.45 | 1.08 | 0.81 | 0.78 | 1.45 | 2.83 | 1.71 | 0 | 0.24 | 0.64 | 2.39 | 2.7 | 0.34 | 0.44 | 3.03 | 0 | 0.69 | 0 | 1.21 | 3.37 | 0.45 | 0 | 0.07 | 0.67 | 1.01
| Analytical results for Pd in OREAS 600 (Indicative Value < 5 ppb) | [29 empty] | Indicative
Pd | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 22; 01: 25 | ppb
|  |  | 00: FA*OES; 01: FA*OES | 0
|  |  | 00: 30g; 01: 40g | 0
| 1 | 1 | 01: < 5 | 1
| 1 | 2 | 01: < 5 | 9
| 1 | 3 | 01: < 5 | 16
| 1 | 4 | 00: < 100; 01: < 5 | < 5
| 1 | 5 | 00: < 100; 01: < 5 | 7
| 1 | 6 | 00: < 100; 01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00 to 01
| Analytical results for Pt in OREAS 600 (Indicative Value < 5 ppb) | [29 empty] | Indicative
Pt | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 22; 01: 25 | ppb
|  |  | 00: FA*OES; 01: FA*OES | 0
|  |  | 00: 30g; 01: 40g | 0
| 1 | 1 | 01: < 5 | 1
| 1 | 2 | 01: < 5 | 6
| 1 | 3 | 01: 5 | 16
| 1 | 4 | 00: < 100; 01: < 5 | < 5
| 1 | 5 | 00: 100; 01: < 5 | 7
| 1 | 6 | 00: < 100; 01: < 5
| Mean |  | 00: 100; 01: 5
| Median |  | 00: 100; 01: 5
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 00: 0.67; 01: 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on D21: The INAA data is omitted from determination of the certified value for Au and is used solely for the calculation of Tolerance Limits and homogeneity evaluation of OREAS 600.

## Aqua Regia
| Analytical results for Ag in OREAS 600 (Certified Value 24.3 ppm) | [23 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*AAS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 2
| 1 | 1 | 25.0914 | 23.37 | 24.6 | 24 | 25.5 | 25.9 | 24.8 | 23.2986 | 25.7 | 25.5 | 24.3665 | 23.9 | 23.4 | 24 | > 10 | 21 | 23.983 | 22 | 24.5 | 26 | 20.4 |  | 1
| 1 | 2 | 24.896 | 23.25 | 24.5 | 24.8 | 25.4 | 25.5 | 24.2 | 22.8404 | 25.4 | 24.6 | 24.52 | 23.8 | 24 | 24 | > 10 | 22 | 24.039 | 23 | 23.9 | 28 | 19.6 |  | 1
| 1 | 3 | 25.535 | 22.81 | 24.5 | 25.1 | 25.3 | 25.6 | 24.3 | 22.9633 | 24.8 | 24.3 | 24.2975 | 24.1 | 23.8 | 24 | > 10 | 21 | 24.173 | 22 | 24 | 28 | 20.8 |  | 16
| 1 | 4 | 24.8617 | 22.07 | 24.2 | 24.8 | 25.8 | 25.6 | 24.2 | 23.2223 | 25.5 | 24.7 | 24.8636 | 23.4 | 23.6 | 24 | > 10 | 22 | 24.656 | 23 | 24.2 | 29 | 20.6 |  | 24.3118
| 1 | 5 | 25.1429 | 22.99 | 24.9 | 24.4 | 25.7 | 25.3 | 24.2 | 23.0211 | 24.9 | 24.7 | 25.4867 | 24.6 | 24 | 24 | > 10 | 22 | 24.467 | 23 | 23.8 | 27 | 20.7 |  | 66
| 1 | 6 | 25.25 | 23.73 | 24.4 | 24.7 | 25.2 | 26.3 | 24.2 | 23.4468 | 25.3 | 24.5 | 25.3357 | 23.6 | 23.7 | 24 | > 10 | 22 | 25.067 | 23 | 23.6 | 29 | 21.2
| Mean |  | 25.1295 | 23.0367 | 24.5167 | 24.6333 | 25.4833 | 25.7 | 24.3167 | 23.1321 | 25.2667 | 24.7167 | 24.8117 | 23.9 | 23.75 | 24 |  | 21.6667 | 24.3975 | 22.6667 | 24 | 27.8333 | 20.55
| Median |  | 25.1171 | 23.12 | 24.5 | 24.75 | 25.45 | 25.6 | 24.2 | 23.1217 | 25.35 | 24.65 | 24.6918 | 23.85 | 23.75 | 24 |  | 22 | 24.32 | 23 | 23.95 | 28 | 20.65
| Std Dev. |  | 0.24772 | 0.57015 | 0.23166 | 0.38297 | 0.23166 | 0.35214 | 0.24014 | 0.22837 | 0.35024 | 0.41191 | 0.50608 | 0.41952 | 0.23452 | 0 |  | 0.5164 | 0.41728 | 0.5164 | 0.31623 | 1.169 | 0.53572
| Rel.Std.Dev. |  | 0.9858% | 2.475% | 0.9449% | 1.555% | 0.9091% | 1.37% | 0.9875% | 0.9873% | 1.386% | 1.667% | 2.04% | 1.755% | 0.9875% | 0% |  | 2.383% | 1.71% | 2.278% | 1.318% | 4.2% | 2.607%
| PDM3 |  | 3.363% | -5.245% | 0.8426% | 1.323% | 4.819% | 5.71% | 0.01998% | -4.852% | 3.928% | 1.665% | 2.056% | -1.694% | -2.311% | -1.283% |  | -10.88% | 0.3525% | -6.767% | -1.283% | 14.48% | -15.47%
| Z-Score (Absolute) |  | 0.76 | 1.29 | 0.16 | 0.27 | 1.1 | 1.31 | 0.04 | 1.2 | 0.89 | 0.2 | 0.44 | 0.45 | 0.59 | 0.35 | NA | 2.63 | 0.04 | 1.65 | 0.35 | 3.4 | 3.72
| Analytical results for Al in OREAS 600 (Certified Value 0.984 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 1.67 | 1.6224 | 1.08 | 1 | 1.01 | 0.99 | 1.5651 | 0.65 | 0.94 | 1.17 | 1.1974 | 1.008 | 1.45 | 0.96 | 1.19 | 0.75 | 0.922 | 0.7 | 0.95 | 1.16 | 0.99 |  | 1
| 1 | 2 | 1.6667 | 1.5608 | 1.12 | 1.01 | 1.02 | 0.97 | 1.6284 | 0.78 | 0.93 | 1.11 | 1.1814 | 0.995 | 1.35 | 0.97 | 1.21 | 0.78 | 0.93 | 0.71 | 0.93 | 1.06 | 0.97 |  | 0
| 1 | 3 | 1.66 | 1.5602 | 1.09 | 1 | 1.02 | 0.99 | 1.6149 | 0.74 | 0.95 | 1.11 | 1.2322 | 1.037 | 1.35 | 0.97 | 1.07 | 0.76 | 0.935 | 0.72 | 0.9 | 1.11 | 0.99 |  | 16
| 1 | 4 | 1.6567 | 1.5095 | 1.13 | 1.01 | 1.02 | 0.98 | 1.4782 | 0.64 | 0.93 | 1.13 | 1.2966 | 0.999 | 1.38 | 0.99 | 1.2 | 0.76 | 0.926 | 0.74 | 0.92 | 1.11 | 0.95 |  | 0.98431
| 1 | 5 | 1.6633 | 1.5412 | 1.09 | 1 | 1.01 | 0.96 | 1.5959 | 0.69 | 0.95 | 1.13 | 1.2784 | 1.004 | 1.39 | 1.01 | 1.21 | 0.75 | 0.942 | 0.77 | 0.93 | 0.98 | 0.95 |  | 67
| 1 | 6 | 1.6567 | 1.5658 | 1.11 | 0.98 | 1.02 | 0.95 | 1.7482 | 0.81 | 0.94 | 1.15 | 1.2429 | 1.02 | 1.4 | 1.03 | 1.21 | 0.75 | 0.919 | 0.77 | 0.9 | 1.15 | 0.96
| Mean |  | 1.6622 | 1.56 | 1.1033 | 1 | 1.0167 | 0.97333 | 1.6051 | 0.71833 | 0.94 | 1.1333 | 1.2382 | 1.0105 | 1.3867 | 0.98833 | 1.1817 | 0.75833 | 0.929 | 0.735 | 0.92167 | 1.095 | 0.96833
| Median |  | 1.6617 | 1.5605 | 1.1 | 1 | 1.02 | 0.975 | 1.6054 | 0.715 | 0.94 | 1.13 | 1.2376 | 1.006 | 1.385 | 0.98 | 1.205 | 0.755 | 0.928 | 0.73 | 0.925 | 1.11 | 0.965
| Std Dev. |  | 0.0054433 | 0.036964 | 0.019664 | 0.010954 | 0.005164 | 0.01633 | 0.08821 | 0.06969 | 0.0089443 | 0.023381 | 0.044647 | 0.015579 | 0.037238 | 0.027142 | 0.055287 | 0.01169 | 0.0085323 | 0.030166 | 0.019408 | 0.066558 | 0.018348
| Rel.Std.Dev. |  | 0.3275% | 2.37% | 1.782% | 1.095% | 0.5079% | 1.678% | 5.496% | 9.702% | 0.9515% | 2.063% | 3.606% | 1.542% | 2.685% | 2.746% | 4.679% | 1.542% | 0.9184% | 4.104% | 2.106% | 6.078% | 1.895%
| PDM3 |  | 68.87% | 58.48% | 12.09% | 1.594% | 3.287% | -1.116% | 63.07% | -27.02% | -4.502% | 15.14% | 25.79% | 2.66% | 40.88% | 0.4084% | 20.05% | -22.96% | -5.619% | -25.33% | -6.365% | 11.25% | -1.623%
| Z-Score (Absolute) |  | 4.73 | 3.99 | 0.67 | 0.08 | 0.04 | 0.27 | 4.32 | 2.12 | 0.51 | 0.89 | 1.65 | 0 | 2.73 | 0.16 | 1.41 | 1.83 | 0.59 | 2 | 0.65 | 0.61 | 0.31
| Analytical results for As in OREAS 600 (Certified Value 85 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 1
| 1 | 1 | 93.01 | 87 | 85 | 90 | 91 | 89 | 88.7518 | 90 | 90 | 81 | 78.9721 | 88 | 97.41 | 79 | 67 | 67 | 85.341 | 80 | 90 | 74.5 | 93 |  | 1
| 1 | 2 | 93.13 | 84 | 92 | 91 | 89 | 88 | 88.0136 | 90 | 91 | 81 | 79.3358 | 88 | 91.58 | 84 | 70 | 76 | 86.174 | 82 | 89 | 81.5 | 95 |  | 0
| 1 | 3 | 92.88 | 85 | 94 | 91 | 90 | 88 | 87.2019 | 80 | 89 | 76 | 82.0656 | 88 | 91.92 | 81 | 59 | 72 | 87.709 | 78 | 88 | 81 | 95 |  | 16
| 1 | 4 | 92.195 | 82 | 94 | 93 | 94 | 85 | 88.4272 | 70 | 90 | 77 | 80.9843 | 88 | 93.64 | 72 | 75 | 66 | 84.048 | 75 | 87 | 80.5 | 91.6 |  | 85.3495
| 1 | 5 | 93.115 | 82 | 89 | 93 | 92 | 89 | 90.3097 | 90 | 89 | 78 | 79.4306 | 88 | 93.31 | 83 | 67 | 75 | 83.534 | 80 | 87 | 80.5 | 93 |  | 68
| 1 | 6 | 93.295 | 89 | 90 | 92 | 95 | 85 | 89.2527 | 90 | 91 | 80 | 80.8186 | 86 | 90.99 | 81 | 72 | 67 | 84.198 | 85 | 86 | 79.5 | 94
| Mean |  | 92.9375 | 84.8333 | 90.6667 | 91.6667 | 91.8333 | 87.3333 | 88.6595 | 85 | 90 | 78.8333 | 80.2679 | 87.6667 | 93.1417 | 80 | 68.3333 | 70.5 | 85.1673 | 80 | 87.8333 | 79.5833 | 93.6
| Median |  | 93.0625 | 84.5 | 91 | 91.5 | 91.5 | 88 | 88.5895 | 90 | 90 | 79 | 80.1246 | 88 | 92.615 | 81 | 68.5 | 69.5 | 84.7695 | 80 | 87.5 | 80.5 | 93.5
| Std Dev. |  | 0.38891 | 2.7869 | 3.4448 | 1.2111 | 2.3166 | 1.8619 | 1.0658 | 8.3666 | 0.89443 | 2.137 | 1.2081 | 0.8165 | 2.3262 | 4.2895 | 5.5015 | 4.4159 | 1.5725 | 3.4059 | 1.472 | 2.5771 | 1.3266
| Rel.Std.Dev. |  | 0.4185% | 3.285% | 3.799% | 1.321% | 2.523% | 2.132% | 1.202% | 9.843% | 0.9938% | 2.711% | 1.505% | 0.9314% | 2.497% | 5.362% | 8.051% | 6.264% | 1.846% | 4.257% | 1.676% | 3.238% | 1.417%
| PDM3 |  | 8.89% | -0.6048% | 6.23% | 7.401% | 7.597% | 2.324% | 3.878% | -0.4095% | 5.449% | -7.635% | -5.954% | 2.715% | 9.13% | -6.268% | -19.94% | -17.4% | -0.2135% | -6.268% | 2.91% | -6.756% | 9.667%
| Z-Score (Absolute) |  | 0.84 | 0.37 | 0.5 | 0.65 | 0.67 | 0 | 0.2 | 0.35 | 0.4 | 1.27 | 1.06 | 0.05 | 0.87 | 0.86 | 2.57 | 2.52 | 0.32 | 1.1 | 0.07 | 1.01 | 0.94
| Analytical results for Au in OREAS 600 (Certified Value 0.192 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 00 | 01 | 02 | 05 | 06 | 07 | 09 | 14 | 15 | 18 | 24 | 25 | 26 | 27 | [8 empty] | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | [8 empty] | 3
|  |  | 1.0g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 50g | 40g | 40g | 40g | 40g | [8 empty] | 3
| 1 | 1 | 0.193 | 0.198 | 0.173 | 0.2 | 0.16 | 0.2 | 0.17463 | 0.181 | 0.18 | 0.193 | 0.192 | 0.197 | 0.21 | 0.2595 | [8 empty] | 1
| 1 | 2 | 0.199 | 0.198 | 0.174 | 0.2 | 0.2 | 0.2 | 0.17154 | 0.194 | 0.19 | 0.198 | 0.205 | 0.192 | 0.19 | 0.2195 | [8 empty] | 0
| 1 | 3 | 0.207 | 0.205 | 0.182 | 0.19 | 0.19 | 0.2 | 0.17213 | 0.183 | 0.2 | 0.203 | 0.192 | 0.194 | 0.2 | 0.229 | [8 empty] | 16
| 1 | 4 | 0.193 | 0.204 | 0.174 | 0.2 | 0.2 | 0.2 | 0.18078 | 0.186 | 0.19 | 0.203 | 0.199 | 0.192 | 0.19 | 0.227 | [8 empty] | 0.19246
| 1 | 5 | 0.209 | 0.192 | 0.181 | 0.21 | 0.18 | 0.2 | 0.18295 | 0.177 | 0.2 | 0.197 | 0.201 | 0.193 | 0.19 | 0.247 | [8 empty] | 69
| 1 | 6 | 0.195 | 0.198 | 0.177 | 0.2 | 0.22 | 0.2 | 0.18495 | 0.18 | 0.2 | 0.198 | 0.194 | 0.2 | 0.2 | 0.2365
|  | 7 | 01: 0.203
|  | 8 | 01: 0.195
|  | 9 | 01: 0.191
|  | 10 | 01: 0.199
|  | 11 | 01: 0.201
|  | 12 | 01: 0.209
|  | 13 | 01: 0.19
|  | 14 | 01: 0.202
|  | 15 | 01: 0.201
|  | 16 | 01: 0.186
|  | 17 | 01: 0.198
|  | 18 | 01: 0.2
|  | 19 | 01: 0.198
|  | 20 | 01: 0.197
| Mean |  | 0.1983 | 0.19917 | 0.17683 | 0.2 | 0.19167 | 0.2 | 0.17783 | 0.1835 | 0.19333 | 0.19867 | 0.19717 | 0.19467 | 0.19667 | 0.23642
| Median |  | 0.1985 | 0.198 | 0.1755 | 0.2 | 0.195 | 0.2 | 0.17771 | 0.182 | 0.195 | 0.198 | 0.1965 | 0.1935 | 0.195 | 0.23275
| Std Dev. |  | 0.0061224 | 0.0047504 | 0.0038687 | 0.0063246 | 0.020412 | 3.0405e-17 | 0.0057938 | 0.0059582 | 0.008165 | 0.0038297 | 0.0053448 | 0.0032042 | 0.008165 | 0.014661
| Rel.Std.Dev. |  | 3.087% | 2.385% | 2.188% | 3.162% | 10.65% | 1.52e-14% | 3.258% | 3.247% | 4.223% | 1.928% | 2.711% | 1.646% | 4.152% | 6.201%
| PDM3 |  | 3.035% | 3.486% | -8.118% | 3.919% | -0.4112% | 3.919% | -7.601% | -4.655% | 0.4548% | 3.226% | 2.447% | 1.148% | 2.187% | 22.84%
| Z-Score (Absolute) |  | NA | 0.51 | 4.01 | 0.67 | 1.01 | 0.67 | 3.81 | 2.66 | 0.67 | 0.4 | 0.1 | 0.4 | 0 | 8.04
| Analytical results for B in OREAS 600 (Certified Value < 10 ppm) | [23 empty] | Certified
B | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 05; 02: 06; 05: 07; 06: 08; 07: 09; 08: 14; 09: 18; 10: 19; 11: 21; 12: 23 |  | ppm
|  |  | 01: AR*OES; 02: AR*OES; 05: AR*OES; 06: AR*OES; 07: AR*OES/MS; 08: AR*OES; 09: AR*OES; 10: AR*OES/MS; 11: AR*OES; 12: AR*OES |  | 1
|  |  | 01: 01g; 02: 01g; 05: 01g; 06: 01g; 07: 20g; 08: 0.5g; 09: 10g; 10: 0.25g; 11: 0.25g; 12: 0.5g |  | 1
| 1 | 1 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 14.604; 11: 12; 12: < 20 |  | 1
| 1 | 2 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 14.568; 11: 16; 12: < 20 |  | 3
| 1 | 3 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 14.016; 11: 15; 12: < 20 |  | 16
| 1 | 4 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 14.202; 11: 12; 12: < 20 |  | < 10
| 1 | 5 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 14.727; 11: 12; 12: < 20 |  | 70
| 1 | 6 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 14.091; 11: 12; 12: < 20
| Mean |  | 10: 14.368; 11: 13.1667
| Median |  | 10: 14.385; 11: 12
| Std Dev. |  | 10: 0.30092; 11: 1.8348
| Rel.Std.Dev. |  | 10: 2.094%; 11: 13.94%
PDM3
| Z-Score (Absolute) |  | 01: NA; 02: NA; 05: NA; 06: NA; 07: NA; 08: NA; 09: NA; 10: 0.67; 11: 0.67; 12: NA
| Analytical results for Ba in OREAS 600 (Indicative Value 1289 ppm) | [23 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 23 | 25 | [5 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | [5 empty] | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g | [5 empty] | 0
| 1 | 1 | 34.6267 | 1208 | 190 | 40 | 60 | 190 | 179.8493 | 1325 | 296 | 470.374 | 110 | 286 | 138 | 237 | 807 | 25 | 60 | [5 empty] | 1
| 1 | 2 | 34.1813 | 1277 | 210 | 40 | 50 | 190 | 182 | 1338 | 319 | 469.998 | 102 | 288 | 139 | 261 | 804 | 27 | 65 | [5 empty] | 4
| 1 | 3 | 34.5973 | 1208 | 200 | 40 | 40 | 200 | 176.1029 | 1408 | 328 | 482.794 | 96 | 297 | 143 | 266 | 831 | 28 | 63 | [5 empty] | 16
| 1 | 4 | 34.4187 | 1199 | 200 | 40 | 40 | 180 | 152.6925 | 1336 | 342 | 482.374 | 90 | 243 | 143 | 259 | 816 | 25 | 71 | [5 empty] | 1288.75
| 1 | 5 | 34.8907 | 1151 | 200 | 40 | 50 | 180 | 208.6531 | 1363 | 330 | 475.67 | 87 | 233 | 138 | 263 | 884 | 28 | 73 | [5 empty] | 12
| 1 | 6 | 35.072 | 1241 | 200 | 30 | 60 | 180 | 202 | 1411 | 322 | 481.259 | 84 | 310 | 144 | 252 | 844 | 28 | 77
| Mean |  | 34.6311 | 1214 | 200 | 38.3333 | 50 | 186.6667 | 183.5496 | 1363.5 | 322.8333 | 477.0782 | 94.8333 | 276.1667 | 140.8333 | 256.3333 | 831 | 26.8333 | 68.1667
| Median |  | 34.612 | 1208 | 200 | 40 | 50 | 185 | 180.9246 | 1350.5 | 325 | 478.4645 | 93 | 287 | 141 | 260 | 823.5 | 27.5 | 68
| Std Dev. |  | 0.31948 | 42.3415 | 6.3246 | 4.0825 | 8.9443 | 8.165 | 19.9715 | 37.7505 | 15.3677 | 5.9204 | 9.8472 | 30.9155 | 2.7869 | 10.5767 | 30.0267 | 1.472 | 6.5243
| Rel.Std.Dev. |  | 0.9225% | 3.488% | 3.162% | 10.65% | 17.89% | 4.374% | 10.88% | 2.769% | 4.76% | 1.241% | 10.38% | 11.19% | 1.979% | 4.126% | 3.613% | 5.486% | 9.571%
| PDM3 |  | -97.31% | -5.8% | -84.48% | -97.03% | -96.12% | -85.52% | -85.76% | 5.8% | -74.95% | -62.98% | -92.64% | -78.57% | -89.07% | -80.11% | -35.52% | -97.92% | -94.71%
| Z-Score (Absolute) |  | 0.75 | 5.09 | 0.07 | 0.73 | 0.68 | 0 | 0.02 | 5.83 | 0.67 | 1.44 | 0.45 | 0.44 | 0.23 | 0.36 | 3.19 | 0.79 | 0.59
| Analytical results for Be in OREAS 600 (Certified Value 0.7 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | [6 empty] | 3
| 1 | 1 | < 1 | 0.62 | 0.7 | 0.7 | 0.7 | 0.7 | 0.72313 | 0.6 | 0.71465 | 0.87 | 0.83 | 0.7 | 0.5 | < 1 | 0.62 | < 2 | [6 empty] | 1
| 1 | 2 | < 1 | 0.65 | 0.7 | 0.7 | 0.7 | 0.7 | 0.76452 | 0.6 | 0.7 | 0.81 | 0.75 | 0.7 | 0.4 | < 1 | 0.65 | < 2 | [6 empty] | 0
| 1 | 3 | < 1 | 0.65 | 0.7 | 0.7 | 0.7 | 0.7 | 0.72346 | 0.6 | 0.69999 | 0.84 | 0.8 | 0.7 | 0.4 | < 1 | 0.7 | < 2 | [6 empty] | 16
| 1 | 4 | < 1 | 0.64 | 0.7 | 0.7 | 0.7 | 0.7 | 0.73089 | 0.6 | 0.67308 | 0.75 | 0.8 | 0.7 | 0.4 | < 1 | 0.71 | < 2 | [6 empty] | 0.70055
| 1 | 5 | < 1 | 0.63 | 0.7 | 0.7 | 0.7 | 0.7 | 0.73003 | 0.6 | 0.67468 | 0.8 | 0.81 | 0.7 | 0.4 | < 1 | 0.66 | < 2 | [6 empty] | 71
| 1 | 6 | < 1 | 0.66 | 0.7 | 0.7 | 0.7 | 0.7 | 0.73145 | 0.6 | 0.66063 | 0.64 | 0.78 | 0.7 | 0.5 | < 1 | 0.68 | < 2
| Mean |  |  | 0.64167 | 0.7 | 0.7 | 0.7 | 0.7 | 0.73391 | 0.6 | 0.68717 | 0.785 | 0.795 | 0.7 | 0.43333 |  | 0.67
| Median |  |  | 0.645 | 0.7 | 0.7 | 0.7 | 0.7 | 0.73046 | 0.6 | 0.68734 | 0.805 | 0.8 | 0.7 | 0.4 |  | 0.67
| Std Dev. |  |  | 0.01472 | 1.2162e-16 | 1.2162e-16 | 1.2162e-16 | 1.2162e-16 | 0.015443 | 0 | 0.020701 | 0.08167 | 0.027386 | 1.2162e-16 | 0.05164 |  | 0.033466
| Rel.Std.Dev. |  |  | 2.294% | 1.737e-14% | 1.737e-14% | 1.737e-14% | 1.737e-14% | 2.104% | 0% | 3.012% | 10.4% | 3.445% | 1.737e-14% | 11.92% |  | 4.995%
| PDM3 |  |  | -8.406% | -0.07888% | -0.07888% | -0.07888% | -0.07888% | 4.762% | -14.35% | -1.91% | 12.05% | 13.48% | -0.07888% | -38.14% |  | -4.361%
| Z-Score (Absolute) |  | NA | 1.42 | 0 | 0 | 0 | 0 | 0.67 | 2.43 | 0.31 | 2.06 | 2.3 | 0 | 6.47 | NA | 0.73 | NA
| Analytical results for Bi in OREAS 600 (Certified Value 6.09 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 2
| 1 | 1 | 6.935 | 5.92 | 4 | 10 | 7 | 6 | 6.7047 | < 10 | 6 | 3 | 4.8783 | 6 | 7.56 | < 10 | 4.58 | < 10 | 6.12 | 3 | 6 | 5.7 | 6.4 |  | 1
| 1 | 2 | 7.2767 | 6.23 | 5 | 11 | 8 | 7 | 6.1617 | 10 | 5 | 4 | 4.7502 | 6 | 6.88 | < 10 | 4.62 | < 10 | 6.243 | 3 | 7 | 5.9 | 6.48 |  | 0
| 1 | 3 | 7.0533 | 6.42 | 5 | 9 | 8 | 7 | 6.7792 | 10 | 5 | 8 | 5.1 | 6 | 7.21 | < 10 | 3.8 | 11 | 6.459 | 4 | 6 | 6.1 | 6.6 |  | 16
| 1 | 4 | 7.0233 | 6.29 | 4 | 9 | 7 | 6 | 6.4582 | 10 | 5 | 3 | 4.5863 | 6 | 7.34 | < 10 | 5.05 | 11 | 6.51 | 3 | 5 | 6.2 | 6.42 |  | 6.0949
| 1 | 5 | 7.1133 | 6.3 | 4 | 8 | 8 | 6 | 6.6842 | < 10 | 8 | 8 | 4.9068 | 6 | 7.23 | < 10 | 4.39 | 11 | 6.384 | 3 | 6 | 6.4 | 6.54 |  | 72
| 1 | 6 | 7.2 | 6.41 | 4 | 10 | 6 | 6 | 6.3183 | 10 | 8 | 8 | 4.9244 | 6 | 7.28 | < 10 | 4.54 | < 10 | 6.482 | 3 | 6 | 6.3 | 6.72
| Mean |  | 7.1003 | 6.2617 | 4.3333 | 9.5 | 7.3333 | 6.3333 | 6.5177 | 10 | 6.1667 | 5.6667 | 4.8577 | 6 | 7.25 |  | 4.4967 | 11 | 6.3663 | 3.1667 | 6 | 6.1 | 6.5267
| Median |  | 7.0833 | 6.295 | 4 | 9.5 | 7.5 | 6 | 6.5712 | 10 | 5.5 | 6 | 4.8925 | 6 | 7.255 |  | 4.56 | 11 | 6.4215 | 3 | 6 | 6.15 | 6.51
| Std Dev. |  | 0.12383 | 0.1828 | 0.5164 | 1.0488 | 0.8165 | 0.5164 | 0.2454 | 0 | 1.472 | 2.582 | 0.17391 | 0 | 0.22091 |  | 0.40668 | 0 | 0.15415 | 0.40825 | 0.63246 | 0.26077 | 0.12044
| Rel.Std.Dev. |  | 1.744% | 2.919% | 11.92% | 11.04% | 11.13% | 8.154% | 3.765% | 0% | 23.87% | 45.56% | 3.58% | 0% | 3.047% |  | 9.044% | 0% | 2.421% | 12.89% | 10.54% | 4.275% | 1.845%
| PDM3 |  | 16.5% | 2.737% | -28.9% | 55.87% | 20.32% | 3.912% | 6.938% | 64.07% | 1.178% | -7.026% | -20.3% | -1.557% | 18.95% |  | -26.22% | 80.48% | 4.454% | -48.04% | -1.557% | 0.08408% | 7.084%
| Z-Score (Absolute) |  | 0.72 | 0 | 1.88 | 2.98 | 0.94 | 0 | 0.18 | 3.45 | 0.16 | 0.63 | 1.39 | 0.31 | 0.86 | NA | 1.6 | 4.39 | 0.03 | 2.98 | 0.31 | 0.22 | 0.18
| Analytical results for Ca in OREAS 600 (Certified Value 1.76 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 1.54 | 1.78 | 1.71 | 1.85 | 1.88 | 1.84 | 1.6558 | 1.93 | 1.85 | 1.97 | 1.4716 | 1.77 | 1.81 | 1.59 | 1.8 | 1.58 | 1.764 | 1.68 | 1.74 | 1.79 | 1.86 |  | 1
| 1 | 2 | 1.5367 | 1.79 | 1.78 | 1.84 | 1.86 | 1.8 | 1.6909 | 2 | 1.81 | 1.88 | 1.5224 | 1.73 | 1.82 | 1.63 | 1.79 | 1.61 | 1.763 | 1.71 | 1.75 | 1.79 | 1.86 |  | 0
| 1 | 3 | 1.5233 | 1.8 | 1.75 | 1.84 | 1.86 | 1.85 | 1.5928 | 2.02 | 1.83 | 1.87 | 1.5124 | 1.77 | 1.82 | 1.59 | 1.63 | 1.58 | 1.776 | 1.69 | 1.68 | 1.8 | 1.85 |  | 16
| 1 | 4 | 1.5233 | 1.81 | 1.78 | 1.86 | 1.88 | 1.8 | 1.6337 | 1.76 | 1.85 | 1.97 | 1.5523 | 1.75 | 1.82 | 1.65 | 1.83 | 1.61 | 1.76 | 1.73 | 1.71 | 1.79 | 1.85 |  | 1.7567
| 1 | 5 | 1.54 | 1.79 | 1.7 | 1.85 | 1.86 | 1.83 | 1.6595 | 1.94 | 1.8 | 1.91 | 1.5563 | 1.79 | 1.83 | 1.64 | 1.76 | 1.59 | 1.789 | 1.74 | 1.7 | 1.69 | 1.86 |  | 73
| 1 | 6 | 1.5267 | 1.79 | 1.75 | 1.82 | 1.9 | 1.81 | 1.6226 | 1.94 | 1.85 | 1.91 | 1.541 | 1.73 | 1.83 | 1.61 | 1.7 | 1.59 | 1.775 | 1.78 | 1.69 | 1.83 | 1.83
| Mean |  | 1.5317 | 1.7933 | 1.745 | 1.8433 | 1.8733 | 1.8217 | 1.6426 | 1.9317 | 1.8317 | 1.9183 | 1.526 | 1.7567 | 1.8217 | 1.6183 | 1.7517 | 1.5933 | 1.7712 | 1.7217 | 1.7117 | 1.7817 | 1.8517
| Median |  | 1.5317 | 1.79 | 1.75 | 1.845 | 1.87 | 1.82 | 1.6447 | 1.94 | 1.84 | 1.91 | 1.5317 | 1.76 | 1.82 | 1.62 | 1.775 | 1.59 | 1.7695 | 1.72 | 1.705 | 1.79 | 1.855
| Std Dev. |  | 0.0080966 | 0.010328 | 0.033912 | 0.013663 | 0.01633 | 0.02137 | 0.033918 | 0.091742 | 0.022286 | 0.043089 | 0.031604 | 0.024221 | 0.0075277 | 0.025626 | 0.07414 | 0.013663 | 0.010944 | 0.03656 | 0.027869 | 0.047504 | 0.01169
| Rel.Std.Dev. |  | 0.5286% | 0.5759% | 1.943% | 0.7412% | 0.8717% | 1.173% | 2.065% | 4.749% | 1.217% | 2.246% | 2.071% | 1.379% | 0.4132% | 1.583% | 4.233% | 0.8575% | 0.6179% | 2.124% | 1.628% | 2.666% | 0.6313%
| PDM3 |  | -12.81% | 2.085% | -0.6661% | 4.931% | 6.639% | 3.698% | -6.497% | 9.96% | 4.267% | 9.201% | -13.13% | -0.002021% | 3.698% | -7.877% | -0.2866% | -9.3% | 0.8234% | -1.994% | -2.564% | 1.421% | 5.406%
| Z-Score (Absolute) |  | 2.67 | 0.25 | 0.29 | 0.8 | 1.14 | 0.56 | 1.43 | 2.17 | 0.67 | 1.64 | 2.73 | 0.16 | 0.56 | 1.7 | 0.22 | 1.98 | 0 | 0.55 | 0.66 | 0.32 | 0.9
| Analytical results for Cd in OREAS 600 (Certified Value 3.5 ppm) | [23 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 4.1079 | 3.36 | 3.3 | 3.5 | 3.3 | 3.4 | 3.8915 | 6 | 3 | 3.8 | 3.0517 | 3.6 | 3.94 | 3 | 3.47 | 2 | 3.267 | 5 | 3.2 | 3.5 | 3.4 |  | 1
| 1 | 2 | 4.2054 | 3.54 | 3.5 | 3.6 | 3.4 | 3.3 | 3.7575 | 6 | 3 | 3.7 | 3.0834 | 3.7 | 3.76 | 3 | 3.35 | 2 | 3.378 | 5 | 3.2 | 4 | 3.6 |  | 0
| 1 | 3 | 4.1739 | 3.61 | 3.7 | 3.5 | 3.5 | 3.5 | 3.8391 | 6 | 3 | 3.6 | 3.1006 | 3.6 | 3.67 | 3 | 2.74 | 2 | 3.4 | 4 | 3.3 | 3.8 | 3.6 |  | 16
| 1 | 4 | 4.2369 | 3.46 | 3.5 | 3.6 | 3.5 | 3.3 | 3.8543 | 6 | 3 | 3.8 | 3.0639 | 3.5 | 3.72 | 3 | 3.92 | 2 | 3.438 | 5 | 3.4 | 3.9 | 3.6 |  | 3.4951
| 1 | 5 | 4.2369 | 3.45 | 3.4 | 3.4 | 3.4 | 3.3 | 3.943 | 6 | 3 | 3.8 | 3.1584 | 3.6 | 3.75 | 3 | 3.18 | 2 | 3.475 | 5 | 3.3 | 4 | 3.5 |  | 74
| 1 | 6 | 4.1051 | 3.68 | 3.5 | 3.6 | 3.5 | 3.3 | 3.9473 | 6 | 3 | 3.7 | 3.0593 | 3.4 | 3.61 | 3 | 3.15 | 2 | 3.32 | 5 | 3.3 | 3.9 | 3.5
| Mean |  | 4.1777 | 3.5167 | 3.4833 | 3.5333 | 3.4333 | 3.35 | 3.8721 | 6 | 3 | 3.7333 | 3.0862 | 3.5667 | 3.7417 | 3 | 3.3017 | 2 | 3.3797 | 4.8333 | 3.2833 | 3.85 | 3.5333
| Median |  | 4.1896 | 3.5 | 3.5 | 3.55 | 3.45 | 3.3 | 3.8729 | 6 | 3 | 3.75 | 3.0736 | 3.6 | 3.735 | 3 | 3.265 | 2 | 3.389 | 5 | 3.3 | 3.9 | 3.55
| Std Dev. |  | 0.059904 | 0.11673 | 0.13292 | 0.08165 | 0.08165 | 0.083666 | 0.071539 | 0 | 0 | 0.08165 | 0.0396 | 0.10328 | 0.11197 | 0 | 0.39148 | 0 | 0.076369 | 0.40825 | 0.075277 | 0.18708 | 0.08165
| Rel.Std.Dev. |  | 1.434% | 3.319% | 3.816% | 2.311% | 2.378% | 2.497% | 1.848% | 0% | 0% | 2.187% | 1.283% | 2.896% | 2.992% | 0% | 11.86% | 0% | 2.26% | 8.447% | 2.293% | 4.859% | 2.311%
| PDM3 |  | 19.53% | 0.6162% | -0.3375% | 1.093% | -1.768% | -4.152% | 10.79% | 71.67% | -14.17% | 6.815% | -11.7% | 2.047% | 7.054% | -14.17% | -5.535% | -42.78% | -3.304% | 38.29% | -6.06% | 10.15% | 1.093%
| Z-Score (Absolute) |  | 1.98 | 0 | 0.1 | 0.05 | 0.25 | 0.5 | 1.07 | 7.44 | 1.55 | 0.65 | 1.29 | 0.15 | 0.67 | 1.55 | 0.64 | 4.55 | 0.41 | 3.95 | 0.7 | 1.21 | 0.05
| Analytical results for Ce in OREAS 600 (Certified Value 34.8 ppm) | [23 empty] | Certified
Ce | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*OES; 06: AR*MS; 07: AR*OES; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*OES/MS; 11: AR*MS |  | 1
|  |  | 01: 40g; 02: 20g; 05: 0.15g; 06: 01g; 07: 0.5g; 08: 01g; 09: 0.25g; 10: 0.25g; 11: 40g |  | 2
| 1 | 1 | 01: 33.83; 02: 37.6741; 05: 32; 06: 31.178; 07: 31; 08: 40.42; 09: 36.6; 10: 32.226; 11: 35.3 |  | 1
| 1 | 2 | 01: 32.92; 02: 37.9094; 05: 36; 06: 31.6913; 07: 36; 08: 39.36; 09: 36.2; 10: 32.394; 11: 35.5 |  | 0
| 1 | 3 | 01: 33.98; 02: 37.256; 05: 33; 06: 31.0087; 07: 31; 08: 38.27; 09: 29.5; 10: 32.776; 11: 35.9 |  | 16
| 1 | 4 | 01: 33.45; 02: 37.6444; 05: 34; 06: 31.1558; 07: 33; 08: 39.08; 09: 41.2; 10: 32.137; 11: 34.8 |  | 34.804
| 1 | 5 | 01: 34.92; 02: 38.2223; 05: 33; 06: 31.4914; 07: 38; 08: 39.72; 09: 33.2; 10: 32.203; 11: 35.2 |  | 75
| 1 | 6 | 01: 33.87; 02: 37.8605; 05: 36; 06: 31.2859; 07: 35; 08: 39.15; 09: 36; 10: 31.834; 11: 35.1
| Mean |  | 01: 33.8283; 02: 37.7611; 05: 34; 06: 31.3019; 07: 34; 08: 39.3333; 09: 35.45; 10: 32.2617; 11: 35.3
| Median |  | 01: 33.85; 02: 37.7673; 05: 33.5; 06: 31.232; 07: 34; 08: 39.255; 09: 36.1; 10: 32.2145; 11: 35.25
| Std Dev. |  | 01: 0.66101; 02: 0.32277; 05: 1.6733; 06: 0.24918; 07: 2.8284; 08: 0.71548; 09: 3.8914; 10: 0.31151; 11: 0.37417
| Rel.Std.Dev. |  | 01: 1.954%; 02: 0.8548%; 05: 4.922%; 06: 0.7961%; 07: 8.319%; 08: 1.819%; 09: 10.98%; 10: 0.9656%; 11: 1.06%
| PDM3 |  | 01: -2.803%; 02: 8.496%; 05: -2.31%; 06: -10.06%; 07: -2.31%; 08: 13.01%; 09: 1.856%; 10: -7.305%; 11: 1.425%
| Z-Score (Absolute) |  | 01: 0.08; 02: 1.75; 05: 0; 06: 1.25; 07: 0; 08: 2.48; 09: 0.67; 10: 0.81; 11: 0.6
| Analytical results for Co in OREAS 600 (Certified Value 6.69 ppm) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 5.9445 | 7 | 7 | 7 | 7 | 7 | 6.3357 | 8 | 7 | 7 | 5.8396 | 6 | 6.95 | 6 | 5.7 | 7 | 7.09 | 7 | 6 | 6.6 | 7 |  | 1
| 1 | 2 | 5.931 | 7 | 6 | 7 | 7 | 6 | 6.3472 | 8 | 7 | 7 | 6.0071 | 7 | 6.99 | 6 | 5.8 | 7 | 7.14 | 6 | 6 | 6.8 | 8 |  | 0
| 1 | 3 | 5.8455 | 6.9 | 7 | 7 | 7 | 6 | 6.3326 | 8 | 7 | 7 | 5.8293 | 7 | 6.9 | 6 | 4.8 | 7 | 7.14 | 7 | 6 | 6.6 | 8 |  | 16
| 1 | 4 | 5.8365 | 6.9 | 7 | 7 | 7 | 6 | 6.5113 | 7 | 7 | 7 | 5.7589 | 7 | 6.71 | 6 | 6.7 | 6 | 7 | 7 | 6 | 7 | 9 |  | 6.6857
| 1 | 5 | 6.3 | 6.7 | 7 | 7 | 7 | 6 | 6.5769 | 8 | 6 | 7 | 5.7903 | 7 | 7 | 6 | 5.5 | 6 | 6.93 | 8 | 6 | 6.8 | 7 |  | 76
| 1 | 6 | 6.2685 | 6.9 | 7 | 7 | 7 | 7 | 6.5193 | 8 | 7 | 7 | 5.7065 | 6 | 6.83 | 6 | 5.9 | 7 | 7.12 | 7 | 6 | 6.8 | 7
| Mean |  | 6.021 | 6.9 | 6.8333 | 7 | 7 | 6.3333 | 6.4372 | 7.8333 | 6.8333 | 7 | 5.822 | 6.6667 | 6.8967 | 6 | 5.7333 | 6.6667 | 7.07 | 7 | 6 | 6.7667 | 7.6667
| Median |  | 5.9378 | 6.9 | 7 | 7 | 7 | 6 | 6.4293 | 8 | 7 | 7 | 5.8098 | 7 | 6.925 | 6 | 5.75 | 7 | 7.105 | 7 | 6 | 6.8 | 7.5
| Std Dev. |  | 0.20875 | 0.10954 | 0.40825 | 0 | 0 | 0.5164 | 0.11053 | 0.40825 | 0.40825 | 0 | 0.10292 | 0.5164 | 0.11094 | 0 | 0.61536 | 0.5164 | 0.086255 | 0.63246 | 0 | 0.15055 | 0.8165
| Rel.Std.Dev. |  | 3.467% | 1.588% | 5.974% | 0% | 0% | 8.154% | 1.717% | 5.212% | 5.974% | 0% | 1.768% | 7.746% | 1.609% | 0% | 10.73% | 7.746% | 1.22% | 9.035% | 0% | 2.225% | 10.65%
| PDM3 |  | -9.942% | 3.205% | 2.208% | 4.701% | 4.701% | -5.271% | -3.718% | 17.17% | 2.208% | 4.701% | -12.92% | -0.285% | 3.155% | -10.26% | -14.25% | -0.285% | 5.748% | 4.701% | -10.26% | 1.211% | 14.67%
| Z-Score (Absolute) |  | 3.29 | 0.27 | 0 | 0.67 | 0.67 | 2.02 | 1.6 | 4.05 | 0 | 0.67 | 4.09 | 0.67 | 0.26 | 3.37 | 3.7 | 0.67 | 0.96 | 0.67 | 3.37 | 0.27 | 2.29
| Analytical results for Cr in OREAS 600 (Certified Value 23.4 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 21 | 23 | 24 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g | 40g |  |  |  | 1
| 1 | 1 | 25.3098 | 27 | 21 | 25 | 20 | 25 | 23.986 | 23 | 26 | 24 | 22 | 26.09 | 20 | 23 | 23 | 20 | 22 | 20 | 20 |  |  |  | 1
| 1 | 2 | 24.3423 | 28 | 21 | 26 | 26 | 26 | 23.9791 | 32 | 30 | 25 | 23 | 27.01 | 28 | 22 | 24 | 19 | 21 | 20 | 20 |  |  |  | 0
| 1 | 3 | 24.3466 | 26 | 20 | 23 | 26 | 23 | 22.4917 | 21 | 28 | 25 | 29 | 26.25 | 18 | 21 | 20 | 19 | 26 | 20 | 20 |  |  |  | 16
| 1 | 4 | 23.779 | 30 | 21 | 46 | 25 | 21 | 26.5438 | 21 | 24 | 24 | 27 | 26.27 | 24 | 21 | 19 | 21 | 23 | 20 | 20 |  |  |  | 23.4036
| 1 | 5 | 25.8 | 30 | 21 | 25 | 23 | 20 | 24.3946 | 24 | 26 | 24 | 24 | 28.26 | 21 | 22 | 21 | 20 | 22 | 20 | 20 |  |  |  | 77
| 1 | 6 | 23.0179 | 25 | 24 | 26 | 29 | 23 | 26.1358 | 26 | 25 | 28 | 19 | 27.8 | 21 | 22 | 18 | 21 | 26 | 20 | 20
| Mean |  | 24.4326 | 27.6667 | 21.3333 | 28.5 | 24.8333 | 23 | 24.5885 | 24.5 | 26.5 | 25 | 24 | 26.9467 | 22 | 21.8333 | 20.8333 | 20 | 23.3333 | 20 | 20
| Median |  | 24.3444 | 27.5 | 21 | 25.5 | 25.5 | 23 | 24.1903 | 23.5 | 26 | 24.5 | 23.5 | 26.64 | 21 | 22 | 20.5 | 20 | 22.5 | 20 | 20
| Std Dev. |  | 1.0083 | 2.0656 | 1.3663 | 8.6429 | 3.0605 | 2.2804 | 1.509 | 4.1352 | 2.1679 | 1.5492 | 3.5777 | 0.9093 | 3.5214 | 0.75277 | 2.3166 | 0.89443 | 2.1602 | 0 | 0
| Rel.Std.Dev. |  | 4.127% | 7.466% | 6.404% | 30.33% | 12.32% | 9.915% | 6.137% | 16.88% | 8.181% | 6.197% | 14.91% | 3.374% | 16.01% | 3.448% | 11.12% | 4.472% | 9.258% | 0% | 0%
| PDM3 |  | 4.397% | 18.22% | -8.846% | 21.78% | 6.109% | -1.724% | 5.063% | 4.685% | 13.23% | 6.821% | 2.548% | 15.14% | -5.997% | -6.709% | -10.98% | -14.54% | -0.3001% | -14.54% | -14.54%
| Z-Score (Absolute) |  | 0.15 | 1.24 | 1.08 | 0.34 | 0.28 | 0.34 | 0.2 | 0.17 | 0.84 | 0.13 | 0 | 0.99 | 0.67 | 0.73 | 1.07 | 1.35 | 0.22 | 1.35 | 1.35
| Analytical results for Cs in OREAS 600 (Certified Value 2.68 ppm) | [23 empty] | Certified
Cs | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 10: 24; 11: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS |  |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 10: 40g; 11: 40g |  |  |  | 2
| 1 | 1 | 01: 4.07; 02: 3.515; 05: 3.0161; 06: 2.29; 07: 2.86; 08: 2.89; 09: 2.328; 10: 1.3; 11: 1.9 |  |  |  | 1
| 1 | 2 | 01: 4.01; 02: 3.2661; 05: 3.0174; 06: 2.24; 07: 2.79; 08: 3.12; 09: 2.326; 10: 1.5; 11: 1.92 |  |  |  | 0
| 1 | 3 | 01: 3.97; 02: 3.4905; 05: 3.1762; 06: 2.23; 07: 2.71; 08: 2.63; 09: 2.331; 10: 1.4; 11: 1.88 |  |  |  | 16
| 1 | 4 | 01: 3.98; 02: 3.4537; 05: 3.1797; 06: 2.15; 07: 2.81; 08: 2.78; 09: 2.252; 10: 1.4; 11: 1.92 |  |  |  | 2.6812
| 1 | 5 | 01: 3.89; 02: 3.462; 05: 3.264; 06: 2.13; 07: 2.8; 08: 3; 09: 2.289; 10: 1.4; 11: 1.86 |  |  |  | 78
| 1 | 6 | 01: 3.96; 02: 3.4378; 05: 3.2264; 06: 2.15; 07: 2.89; 08: 3.28; 09: 2.263; 10: 1.4; 11: 1.88
| Mean |  | 01: 3.98; 02: 3.4375; 05: 3.1466; 06: 2.1983; 07: 2.81; 08: 2.95; 09: 2.2982; 10: 1.4; 11: 1.8933
| Median |  | 01: 3.975; 02: 3.4579; 05: 3.1779; 06: 2.19; 07: 2.805; 08: 2.945; 09: 2.3075; 10: 1.4; 11: 1.89
| Std Dev. |  | 01: 0.05933; 02: 0.088376; 05: 0.10571; 06: 0.064005; 07: 0.06229; 08: 0.23461; 09: 0.035199; 10: 0.063246; 11: 0.024221
| Rel.Std.Dev. |  | 01: 1.491%; 02: 2.571%; 05: 3.36%; 06: 2.912%; 07: 2.217%; 08: 7.953%; 09: 1.532%; 10: 4.518%; 11: 1.279%
| PDM3 |  | 01: 48.44%; 02: 28.21%; 05: 17.36%; 06: -18.01%; 07: 4.805%; 08: 10.03%; 09: -14.29%; 10: -47.78%; 11: -29.38%
| Z-Score (Absolute) |  | 01: 1.29; 02: 0.73; 05: 0.37; 06: 0.67; 07: 0; 08: 0.15; 09: 0.56; 10: 1.55; 11: 1.01
| Analytical results for Cu in OREAS 600 (Certified Value 488 ppm) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | AR*OES | AR*OES | AR*AAS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g | 0
| 1 | 1 | 499.86 | 471.9 | 487 | 477 | 515 | 516 | 500 | 463.6481 | 492 | 493 | 534 | 448.4816 | 486 | 502.31 | 442 | 487 | 458 | 496.971 | 480 | 482 | 512 | 497 | 1
| 1 | 2 | 503.01 | 476.3 | 482 | 504 | 511 | 510 | 500 | 466.5417 | 510 | 491 | 512 | 458.3439 | 469 | 476.88 | 453 | 486 | 456 | 511.461 | 486 | 489 | 501 | 497 | 0
| 1 | 3 | 491.575 | 475.9 | 491 | 502 | 508 | 514 | 506 | 467.0562 | 554 | 508 | 500 | 457.0588 | 486 | 479.33 | 450 | 482 | 460 | 517.261 | 480 | 482 | 496 | 495 | 16
| 1 | 4 | 500.96 | 466.3 | 479 | 503 | 518 | 521 | 497 | 460.7497 | 456 | 494 | 505 | 474.6863 | 485 | 492.71 | 449 | 478 | 463 | 508.846 | 482 | 481 | 488 | 490 | 488.0741
| 1 | 5 | 501.57 | 470.2 | 492 | 477 | 516 | 507 | 502 | 461.922 | 708 | 508 | 513 | 464.1885 | 491 | 486.95 | 458 | 473 | 457 | 514.625 | 498 | 481 | 509 | 492 | 79
| 1 | 6 | 492.1 | 474.9 | 487 | 489 | 509 | 524 | 496 | 463.3268 | 530 | 502 | 506 | 479.1411 | 475 | 476.72 | 450 | 463 | 456 | 508.8 | 504 | 476 | 498 | 493
| Mean |  | 498.1792 | 472.5833 | 486.3333 | 492 | 512.8333 | 515.3333 | 500.1667 | 463.8741 | 541.6667 | 499.3333 | 511.6667 | 463.65 | 482 | 485.8167 | 450.3333 | 478.1667 | 458.3333 | 509.6607 | 488.3333 | 481.8333 | 500.6667 | 494
| Median |  | 500.41 | 473.4 | 487 | 495.5 | 513 | 515 | 500 | 463.4874 | 520 | 498 | 509 | 461.2662 | 485.5 | 483.14 | 450 | 480 | 457.5 | 510.1535 | 484 | 481.5 | 499.5 | 494
| Std Dev. |  | 5.0192 | 3.8938 | 5.0465 | 12.8374 | 4.0702 | 6.4395 | 3.6009 | 2.4966 | 88.0492 | 7.6855 | 11.9443 | 11.5217 | 8.2462 | 10.2521 | 5.2409 | 9.0646 | 2.7325 | 7.0439 | 10.2307 | 4.1673 | 8.8015 | 2.8284
| Rel.Std.Dev. |  | 1.008% | 0.8239% | 1.038% | 2.609% | 0.7937% | 1.25% | 0.7199% | 0.5382% | 16.26% | 1.539% | 2.334% | 2.485% | 1.711% | 2.11% | 1.164% | 1.896% | 0.5962% | 1.382% | 2.095% | 0.8649% | 1.758% | 0.5726%
| PDM3 |  | 2.07% | -3.174% | -0.3567% | 0.8044% | 5.073% | 5.585% | 2.478% | -4.958% | 10.98% | 2.307% | 4.834% | -5.004% | -1.245% | -0.4625% | -7.733% | -2.03% | -6.093% | 4.423% | 0.05311% | -1.279% | 2.58% | 1.214%
| Z-Score (Absolute) |  | 0.53 | 1.16 | 0.25 | 0.12 | 1.49 | 1.66 | 0.66 | 1.73 | 0.45 | 0.6 | 1.12 | 1.74 | 0.54 | 0.29 | 2.62 | 0.79 | 2.09 | 1.28 | 0.12 | 0.55 | 0.69 | 0.25
| Analytical results for Dy in OREAS 600 (Certified Value 1.26 ppm) | [23 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 02; 02: 09; 04: 13; 05: 14; 06: 15; 07: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS | 2
|  |  | 01: 40g; 02: 20g; 04: 01g; 05: 0.5g; 06: 01g; 07: 40g | 3
| 1 | 1 | 01: 1.22; 02: 1.3597; 04: 1.2568; 05: 1.21; 06: 1.37; 07: 1.18 | 1
| 1 | 2 | 01: 1.23; 02: 1.3371; 04: 1.2948; 05: 1.21; 06: 1.24; 07: 1.2 | 0
| 1 | 3 | 01: 1.28; 02: 1.304; 04: 1.2733; 05: 1.19; 06: 1.25; 07: 1.19 | 16
| 1 | 4 | 01: 1.29; 02: 1.319; 04: 1.2995; 05: 1.2; 06: 1.31; 07: 1.2 | 1.265
| 1 | 5 | 01: 1.28; 02: 1.308; 04: 1.3595; 05: 1.19; 06: 1.34; 07: 1.18 | 80
| 1 | 6 | 01: 1.31; 02: 1.3363; 04: 1.3004; 05: 1.17; 06: 1.36; 07: 1.14
| Mean |  | 01: 1.2683; 02: 1.3274; 04: 1.2974; 05: 1.195; 06: 1.3117; 07: 1.1817
| Median |  | 01: 1.28; 02: 1.3277; 04: 1.2971; 05: 1.195; 06: 1.325; 07: 1.185
| Std Dev. |  | 01: 0.035449; 02: 0.021055; 04: 0.034945; 05: 0.015166; 06: 0.055648; 07: 0.022286
| Rel.Std.Dev. |  | 01: 2.795%; 02: 1.586%; 04: 2.694%; 05: 1.269%; 06: 4.243%; 07: 1.886%
| PDM3 |  | 01: 0.2672%; 02: 4.933%; 04: 2.562%; 05: -5.53%; 06: 3.693%; 07: -6.584%
| Z-Score (Absolute) |  | 01: 0.27; 02: 0.82; 04: 0.27; 05: 1.62; 06: 0.53; 07: 1.71
| Analytical results for Er in OREAS 600 (Certified Value 0.54 ppm) | [23 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 09; 04: 14; 05: 15; 06: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 40g; 02: 20g; 04: 0.5g; 05: 01g; 06: 40g | 3
| 1 | 1 | 01: 0.54; 02: 0.56736; 04: 0.51; 05: 0.58; 06: 0.51 | 1
| 1 | 2 | 01: 0.52; 02: 0.53626; 04: 0.52; 05: 0.55; 06: 0.5 | 0
| 1 | 3 | 01: 0.54; 02: 0.53544; 04: 0.52; 05: 0.58; 06: 0.51 | 16
| 1 | 4 | 01: 0.54; 02: 0.53538; 04: 0.51; 05: 0.57; 06: 0.5 | 0.53736
| 1 | 5 | 01: 0.55; 02: 0.54533; 04: 0.53; 05: 0.59; 06: 0.51 | 81
| 1 | 6 | 01: 0.54; 02: 0.56098; 04: 0.52; 05: 0.59; 06: 0.51
| Mean |  | 01: 0.53833; 02: 0.54679; 04: 0.51833; 05: 0.57667; 06: 0.50667
| Median |  | 01: 0.54; 02: 0.5408; 04: 0.52; 05: 0.58; 06: 0.51
| Std Dev. |  | 01: 0.0098319; 02: 0.014115; 04: 0.0075277; 05: 0.015055; 06: 0.005164
| Rel.Std.Dev. |  | 01: 1.826%; 02: 2.581%; 04: 1.452%; 05: 2.611%; 06: 1.019%
| PDM3 |  | 01: 0.1814%; 02: 1.756%; 04: -3.541%; 05: 7.315%; 06: -5.712%
| Z-Score (Absolute) |  | 01: 0; 02: 0.29; 04: 0.67; 05: 1.29; 06: 1.07
| Analytical results for Eu in OREAS 600 (Certified Value 0.59 ppm) | [23 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 09; 04: 14; 05: 15; 06: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 40g; 02: 20g; 04: 0.5g; 05: 01g; 06: 40g | 3
| 1 | 1 | 01: 0.59; 02: 0.64319; 04: 0.53; 05: 0.62; 06: 0.558 | 1
| 1 | 2 | 01: 0.58; 02: 0.62619; 04: 0.54; 05: 0.59; 06: 0.576 | 0
| 1 | 3 | 01: 0.59; 02: 0.61534; 04: 0.54; 05: 0.62; 06: 0.574 | 16
| 1 | 4 | 01: 0.6; 02: 0.61977; 04: 0.53; 05: 0.63; 06: 0.544 | 0.58757
| 1 | 5 | 01: 0.6; 02: 0.64712; 04: 0.53; 05: 0.63; 06: 0.564 | 82
| 1 | 6 | 01: 0.6; 02: 0.67159; 04: 0.53; 05: 0.58; 06: 0.558
| Mean |  | 01: 0.59333; 02: 0.6372; 04: 0.53333; 05: 0.61167; 06: 0.56233
| Median |  | 01: 0.595; 02: 0.63469; 04: 0.53; 05: 0.62; 06: 0.561
| Std Dev. |  | 01: 0.008165; 02: 0.021072; 04: 0.005164; 05: 0.02137; 06: 0.011827
| Rel.Std.Dev. |  | 01: 1.376%; 02: 3.307%; 04: 0.9682%; 05: 3.494%; 06: 2.103%
| PDM3 |  | 01: 0.9802%; 02: 8.446%; 04: -9.231%; 05: 4.1%; 06: -4.296%
| Z-Score (Absolute) |  | 01: 0; 02: 0.95; 04: 1.31; 05: 0.4; 06: 0.67
| Analytical results for Fe in OREAS 600 (Certified Value 2.22 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 2.212 | 2.26 | 2.07 | 2.2 | 2.18 | 2.14 | 2.2571 | 2.38 | 2.27 | 2.42 | 2.2121 | 2.22 | 2.26 | 2.07 | 2.33 | 2.01 | 2.359 | 1.98 | 2.15 | 2.41 | 2.31 |  | 1
| 1 | 2 | 2.248 | 2.3 | 2.13 | 2.19 | 2.14 | 2.1 | 2.2187 | 2.48 | 2.16 | 2.31 | 2.2489 | 2.18 | 2.28 | 2.13 | 2.32 | 2.06 | 2.365 | 2.02 | 2.12 | 2.23 | 2.31 |  | 0
| 1 | 3 | 2.21 | 2.32 | 2.12 | 2.19 | 2.15 | 2.12 | 2.2071 | 2.49 | 2.12 | 2.31 | 2.2443 | 2.24 | 2.28 | 2.05 | 2.27 | 2.02 | 2.365 | 2.01 | 2.07 | 2.25 | 2.28 |  | 16
| 1 | 4 | 2.23 | 2.29 | 2.16 | 2.21 | 2.16 | 2.09 | 2.2467 | 2.17 | 2.27 | 2.41 | 2.3325 | 2.22 | 2.27 | 2.12 | 2.36 | 2.05 | 2.352 | 2.05 | 2.09 | 2.21 | 2.27 |  | 2.2196
| 1 | 5 | 2.22 | 2.29 | 2.1 | 2.19 | 2.14 | 2.1 | 2.2758 | 2.42 | 2.27 | 2.33 | 2.2917 | 2.3 | 2.29 | 2.12 | 2.27 | 2.03 | 2.395 | 2.08 | 2.08 | 2.27 | 2.29 |  | 83
| 1 | 6 | 2.21 | 2.31 | 2.12 | 2.18 | 2.18 | 2.11 | 2.2345 | 2.41 | 2.36 | 2.36 | 2.3244 | 2.2 | 2.28 | 2.1 | 2.21 | 2.03 | 2.337 | 2.1 | 2.07 | 2.31 | 2.25
| Mean |  | 2.2217 | 2.295 | 2.1167 | 2.1933 | 2.1583 | 2.11 | 2.24 | 2.3917 | 2.2417 | 2.3567 | 2.2756 | 2.2267 | 2.2767 | 2.0983 | 2.2933 | 2.0333 | 2.3622 | 2.04 | 2.0967 | 2.28 | 2.285
| Median |  | 2.216 | 2.295 | 2.12 | 2.19 | 2.155 | 2.105 | 2.2406 | 2.415 | 2.27 | 2.345 | 2.2703 | 2.22 | 2.28 | 2.11 | 2.295 | 2.03 | 2.362 | 2.035 | 2.085 | 2.26 | 2.285
| Std Dev. |  | 0.015042 | 0.020736 | 0.030111 | 0.010328 | 0.018348 | 0.017889 | 0.025233 | 0.11652 | 0.087044 | 0.048854 | 0.048166 | 0.041312 | 0.010328 | 0.031885 | 0.053914 | 0.018619 | 0.019188 | 0.045166 | 0.032042 | 0.072388 | 0.023452
| Rel.Std.Dev. |  | 0.6771% | 0.9035% | 1.423% | 0.4709% | 0.8501% | 0.8478% | 1.127% | 4.872% | 3.883% | 2.073% | 2.117% | 1.855% | 0.4536% | 1.52% | 2.351% | 0.9157% | 0.8123% | 2.214% | 1.528% | 3.175% | 1.026%
| PDM3 |  | 0.0942% | 3.398% | -4.636% | -1.182% | -2.759% | -4.937% | 0.9192% | 7.753% | 0.9953% | 6.176% | 2.525% | 0.3195% | 2.572% | -5.462% | 3.323% | -8.391% | 6.424% | -8.091% | -5.538% | 2.722% | 2.948%
| Z-Score (Absolute) |  | 0.22 | 0.67 | 1.51 | 0.57 | 1 | 1.59 | 0 | 2.4 | 0.02 | 1.43 | 0.44 | 0.16 | 0.45 | 1.74 | 0.65 | 2.53 | 1.5 | 2.45 | 1.76 | 0.17 | 0.55
| Analytical results for Ga in OREAS 600 (Certified Value 4.16 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 17 | 19 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 40g | 40g | [6 empty] | 2
| 1 | 1 | 4.8504 | 4.63 | < 10 | < 10 | < 10 | < 10 | 4.9929 | < 5 | 4.5448 | 3.19 | 4.5 | 4 | 3.31 | < 5 | 3.4 | 3.4 | [6 empty] | 1
| 1 | 2 | 4.8816 | 4.26 | < 10 | < 10 | < 10 | < 10 | 4.7286 | 5 | 4.5364 | 3.16 | 4.36 | 4.2 | 3.24 | < 5 | 3.7 | 3.6 | [6 empty] | 0
| 1 | 3 | 4.7196 | 4.35 | < 10 | < 10 | < 10 | < 10 | 4.8367 | 5 | 4.7402 | 3.29 | 4.4 | 3.5 | 3.3 | < 5 | 3.6 | 3.4 | [6 empty] | 16
| 1 | 4 | 5.0016 | 4.28 | < 10 | < 10 | < 10 | < 10 | 4.9453 | 5 | 4.8534 | 3.05 | 4.5 | 4 | 3.29 | < 5 | 3.8 | 3.2 | [6 empty] | 4.1565
| 1 | 5 | 5.064 | 4.31 | < 10 | < 10 | < 10 | < 10 | 5.1011 | 5 | 4.8684 | 3.13 | 4.46 | 4.1 | 3.11 | < 5 | 3.4 | 3.2 | [6 empty] | 84
| 1 | 6 | 5.0856 | 4.46 | < 10 | < 10 | < 10 | < 10 | 4.9371 | < 5 | 4.7612 | 3.13 | 4.44 | 4.4 | 3.13 | < 5 | 3.8 | 3.2
| Mean |  | 01: 4.9338; 02: 4.3817; 09: 4.9236; 10: 5; 11: 4.7174; 12: 3.1583; 13: 4.4433; 14: 4.0333; 15: 3.23; 17: 3.6167; 18: 3.3333
| Median |  | 01: 4.9416; 02: 4.33; 09: 4.9412; 10: 5; 11: 4.7507; 12: 3.145; 13: 4.45; 14: 4.05; 15: 3.265; 17: 3.65; 18: 3.3
| Std Dev. |  | 01: 0.14153; 02: 0.14077; 09: 0.12849; 10: 0; 11: 0.14579; 12: 0.079603; 13: 0.055737; 14: 0.30111; 15: 0.088769; 17: 0.18348; 18: 0.1633
| Rel.Std.Dev. |  | 01: 2.869%; 02: 3.213%; 09: 2.61%; 10: 0%; 11: 3.09%; 12: 2.52%; 13: 1.254%; 14: 7.466%; 15: 2.748%; 17: 5.073%; 18: 4.899%
| PDM3 |  | 01: 18.7%; 02: 5.416%; 09: 18.46%; 10: 20.29%; 11: 13.49%; 12: -24.02%; 13: 6.9%; 14: -2.964%; 15: -22.29%; 17: -12.99%; 18: -19.8%
| Z-Score (Absolute) |  | 0.67 | 0 | NA | NA | NA | NA | 0.66 | 0.75 | 0.43 | 1.32 | 0.12 | 0.33 | 1.23 | NA | 0.8 | 1.12
| Analytical results for Gd in OREAS 600 (Certified Value 1.97 ppm) | [23 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  | 3
| 1 | 1 | 01: 2; 02: 2.0121; 05: 1.9592; 06: 1.86; 07: 2.32; 08: 1.75 |  | 1
| 1 | 2 | 01: 2.01; 02: 2.0667; 05: 1.9323; 06: 1.84; 07: 2.09; 08: 1.8 |  | 0
| 1 | 3 | 01: 2; 02: 2.027; 05: 2.0107; 06: 1.89; 07: 2.24; 08: 1.83 |  | 16
| 1 | 4 | 01: 2; 02: 2.0405; 05: 1.9265; 06: 1.87; 07: 2.2; 08: 1.76 |  | 1.9691
| 1 | 5 | 01: 2; 02: 2.0038; 05: 1.9679; 06: 1.84; 07: 2.12; 08: 1.75 |  | 85
| 1 | 6 | 01: 2.03; 02: 2.0125; 05: 2.013; 06: 1.82; 07: 2.2; 08: 1.76
| Mean |  | 01: 2.0067; 02: 2.0271; 05: 1.9683; 06: 1.8533; 07: 2.195; 08: 1.775
| Median |  | 01: 2; 02: 2.0198; 05: 1.9636; 06: 1.85; 07: 2.2; 08: 1.76
| Std Dev. |  | 01: 0.012111; 02: 0.023349; 05: 0.037203; 06: 0.025033; 07: 0.082885; 08: 0.032711
| Rel.Std.Dev. |  | 01: 0.6035%; 02: 1.152%; 05: 1.89%; 06: 1.351%; 07: 3.776%; 08: 1.843%
| PDM3 |  | 01: 1.91%; 02: 2.947%; 05: -0.04014%; 06: -5.877%; 07: 11.47%; 08: -9.855%
| Z-Score (Absolute) |  | 01: 0.15; 02: 0.31; 05: 0.15; 06: 1.04; 07: 1.61; 08: 1.73
| Analytical results for Ge in OREAS 600 (Indicative Value < 0.1 ppm) | [23 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 11; 02: 15; 05: 17; 06: 19 |  | ppm
|  |  | 01: AR*OES; 02: AR*OES/MS; 05: AR*OES/MS; 06: AR*OES/MS |  | 2
|  |  | 01: 0.15g; 02: 01g; 05: 0.25g; 06: 0.25g |  | 2
| 1 | 1 | 01: < 10; 02: < 0.5; 05: < 0.1; 06: < 0.1 |  | 1
| 1 | 2 | 01: < 10; 02: < 0.5; 05: < 0.1; 06: < 0.1 |  | 7
| 1 | 3 | 01: < 10; 02: < 0.5; 05: < 0.1; 06: < 0.1 |  | 16
| 1 | 4 | 01: < 10; 02: < 0.5; 05: < 0.1; 06: < 0.1 |  | < 0.1
| 1 | 5 | 01: < 10; 02: < 0.5; 05: < 0.1; 06: < 0.1 |  | 13
| 1 | 6 | 01: < 10; 02: < 0.5; 05: < 0.1; 06: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 06
| Analytical results for Hf in OREAS 600 (Certified Value 0.22 ppm) | [23 empty] | Certified
Hf | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 10: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 10: 40g |  | 2
| 1 | 1 | 01: 0.31; 02: 0.21047; 05: 0.15556; 06: 0.22; 07: 0.28; 08: 0.07; 09: 0.262; 10: 0.2 |  | 1
| 1 | 2 | 01: 0.4; 02: 0.23589; 05: 0.15526; 06: 0.22; 07: 0.28; 08: 0.07; 09: 0.249; 10: 0.2 |  | 0
| 1 | 3 | 01: 0.4; 02: 0.21318; 05: 0.13384; 06: 0.21; 07: 0.24; 08: < 0.05; 09: 0.246; 10: 0.2 |  | 16
| 1 | 4 | 01: 0.37; 02: 0.21288; 05: 0.13208; 06: 0.2; 07: 0.28; 08: 0.07; 09: 0.252; 10: 0.2 |  | 0.21581
| 1 | 5 | 01: 0.37; 02: 0.2354; 05: 0.15189; 06: 0.2; 07: 0.25; 08: 0.06; 09: 0.268; 10: 0.2 |  | 86
| 1 | 6 | 01: 0.41; 02: 0.24259; 05: 0.12503; 06: 0.2; 07: 0.26; 08: 0.07; 09: 0.248; 10: 0.2
| Mean |  | 01: 0.37667; 02: 0.22507; 05: 0.14228; 06: 0.20833; 07: 0.265; 08: 0.068; 09: 0.25417; 10: 0.2
| Median |  | 01: 0.385; 02: 0.22429; 05: 0.14286; 06: 0.205; 07: 0.27; 08: 0.07; 09: 0.2505; 10: 0.2
| Std Dev. |  | 01: 0.036697; 02: 0.014379; 05: 0.013491; 06: 0.0098319; 07: 0.017607; 08: 0.0044721; 09: 0.0088185; 10: 3.0405e-17
| Rel.Std.Dev. |  | 01: 9.743%; 02: 6.389%; 05: 9.482%; 06: 4.719%; 07: 6.644%; 08: 6.577%; 09: 3.47%; 10: 1.52e-14%
| PDM3 |  | 01: 74.54%; 02: 4.291%; 05: -34.07%; 06: -3.463%; 07: 22.79%; 08: -68.49%; 09: 17.77%; 10: -7.325%
| Z-Score (Absolute) |  | 01: 2.52; 02: 0.13; 05: 1.17; 06: 0.13; 07: 0.76; 08: 2.31; 09: 0.59; 10: 0.26
| Analytical results for Hg in OREAS 600 (Certified Value < 1 ppm) | [23 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 16 | 17 | 19 | 23 | 24 | 25 | [8 empty] | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | [8 empty] | 2
|  |  | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g | [8 empty] | 2
| 1 | 1 | < 1 | 1 | 1 | < 1 | 0.25718 | < 1 | < 3 | < 1 | < 5 | 0.45 | 0.164 | < 1 | < 0.5 | 0.25 | [8 empty] | 1
| 1 | 2 | < 1 | < 1 | 1 | < 1 | 0.25237 | < 1 | < 3 | < 1 | < 5 | 0.44 | 0.154 | < 1 | < 0.5 | 0.25 | [8 empty] | 4
| 1 | 3 | < 1 | 1 | 1 | < 1 | 0.26135 | < 1 | < 3 | < 1 | < 5 | 0.34 | 0.14 | < 1 | < 0.5 | 0.27 | [8 empty] | 16
| 1 | 4 | < 1 | 1 | < 1 | < 1 | 0.25614 | < 1 | < 3 | < 1 | < 5 | 0.46 | 0.21 | < 1 | < 0.5 | 0.29 | [8 empty] | < 1
| 1 | 5 | < 1 | < 1 | 1 | 1 | 0.29008 | < 1 | < 3 | < 1 | < 5 | 0.45 | 0.189 | < 1 | < 0.5 | 0.3 | [8 empty] | 87
| 1 | 6 | < 1 | 1 | 1 | 1 | 0.23928 | < 1 | < 3 | < 1 | < 5 | 0.53 | 0.166 | < 1 | < 0.5 | 0.28
| Mean |  | 02: 1; 05: 1; 06: 1; 07: 0.2594; 12: 0.445; 13: 0.1705; 16: 0.27333
| Median |  | 02: 1; 05: 1; 06: 1; 07: 0.25666; 12: 0.45; 13: 0.165; 16: 0.275
| Std Dev. |  | 02: 0; 05: 0; 06: 0; 07: 0.016822; 12: 0.06091; 13: 0.025169; 16: 0.020656
| Rel.Std.Dev. |  | 02: 0%; 05: 0%; 06: 0%; 07: 6.485%; 12: 13.69%; 13: 14.76%; 16: 7.557%
PDM3
| Z-Score (Absolute) |  | NA | 1.2 | 1.74 | 0.65 | 0.67 | NA | NA | NA | NA | 0 | 0.96 | NA | NA | 0.63
| Analytical results for Ho in OREAS 600 (Certified Value 0.22 ppm) | [23 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  | 3
| 1 | 1 | 01: 0.22; 02: 0.2218; 05: 0.17138; 06: 0.22; 07: 0.24; 08: 0.215 |  | 1
| 1 | 2 | 01: 0.21; 02: 0.21949; 05: 0.17418; 06: 0.22; 07: 0.21; 08: 0.215 |  | 0
| 1 | 3 | 01: 0.22; 02: 0.2141; 05: 0.18126; 06: 0.22; 07: 0.23; 08: 0.215 |  | 16
| 1 | 4 | 01: 0.22; 02: 0.21758; 05: 0.17488; 06: 0.21; 07: 0.23; 08: 0.215 |  | 0.21894
| 1 | 5 | 01: 0.21; 02: 0.21637; 05: 0.20867; 06: 0.22; 07: 0.24; 08: 0.215 |  | 88
| 1 | 6 | 01: 0.22; 02: 0.21383; 05: 0.18133; 06: 0.22; 07: 0.22; 08: 0.21
| Mean |  | 01: 0.21667; 02: 0.21719; 05: 0.18195; 06: 0.21833; 07: 0.22833; 08: 0.21417
| Median |  | 01: 0.22; 02: 0.21697; 05: 0.17807; 06: 0.22; 07: 0.23; 08: 0.215
| Std Dev. |  | 01: 0.005164; 02: 0.0031054; 05: 0.013691; 06: 0.0040825; 07: 0.01169; 08: 0.0020412
| Rel.Std.Dev. |  | 01: 2.383%; 02: 1.43%; 05: 7.524%; 06: 1.87%; 07: 5.12%; 08: 0.9531%
| PDM3 |  | 01: -1.038%; 02: -0.7966%; 05: -16.89%; 06: -0.2766%; 07: 4.291%; 08: -2.18%
| Z-Score (Absolute) |  | 01: 0.09; 02: 0.09; 05: 13.05; 06: 0.45; 07: 3.69; 08: 0.89
| Analytical results for In in OREAS 600 (Certified Value 0.73 ppm) | [23 empty] | Certified
In | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 13; 08: 14; 09: 15; 10: 17; 11: 19; 12: 24; 13: 25 |  | ppm
|  |  | 01: AR*OES; 02: AR*MS; 05: AR*OES/MS; 06: AR*OES; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*OES/MS; 11: AR*OES/MS; 12: AR*MS; 13: AR*MS |  | 2
|  |  | 01: 0.5g; 02: 40g; 05: 20g; 06: 0.15g; 07: 01g; 08: 0.5g; 09: 01g; 10: 0.25g; 11: 0.25g; 12: 40g; 13: 40g |  | 2
| 1 | 1 | 01: < 5; 02: 0.75; 05: 0.68255; 06: < 10; 07: 0.76498; 08: 0.69; 09: 0.8; 10: 0.69; 11: 0.638; 12: 0.5; 13: 0.655 |  | 1
| 1 | 2 | 01: < 5; 02: 0.78; 05: 0.69613; 06: < 10; 07: 0.82226; 08: 0.71; 09: 0.76; 10: 0.73; 11: 0.659; 12: 1; 13: 0.67 |  | 0
| 1 | 3 | 01: < 5; 02: 0.77; 05: 0.65993; 06: < 10; 07: 0.7927; 08: 0.75; 09: 0.78; 10: 0.56; 11: 0.643; 12: 0.5; 13: 0.67 |  | 16
| 1 | 4 | 01: < 5; 02: 0.74; 05: 0.65472; 06: < 10; 07: 0.77723; 08: 0.69; 09: 0.81; 10: 0.73; 11: 0.65; 12: 1; 13: 0.665 |  | 0.7309
| 1 | 5 | 01: < 5; 02: 0.78; 05: 0.68548; 06: < 10; 07: 0.78591; 08: 0.72; 09: 0.82; 10: 0.66; 11: 0.652; 12: 1; 13: 0.68 |  | 89
| 1 | 6 | 01: < 5; 02: 0.76; 05: 0.69451; 06: < 10; 07: 0.81432; 08: 0.68; 09: 0.76; 10: 0.7; 11: 0.614; 12: 1; 13: 0.68
| Mean |  | 02: 0.76333; 05: 0.67889; 07: 0.7929; 08: 0.70667; 09: 0.78833; 10: 0.67833; 11: 0.64267; 12: 0.83333; 13: 0.67
| Median |  | 02: 0.765; 05: 0.68401; 07: 0.78931; 08: 0.7; 09: 0.79; 10: 0.695; 11: 0.6465; 12: 1; 13: 0.67
| Std Dev. |  | 02: 0.01633; 05: 0.017561; 07: 0.021885; 08: 0.02582; 09: 0.025626; 10: 0.063692; 11: 0.01582; 12: 0.2582; 13: 0.0094868
| Rel.Std.Dev. |  | 02: 2.139%; 05: 2.587%; 07: 2.76%; 08: 3.654%; 09: 3.251%; 10: 9.389%; 11: 2.462%; 12: 30.98%; 13: 1.416%
| PDM3 |  | 02: 4.437%; 05: -7.117%; 07: 8.482%; 08: -3.316%; 09: 7.858%; 10: -7.192%; 11: -12.07%; 12: 14.01%; 13: -8.332%
| Z-Score (Absolute) |  | 01: NA; 02: 0.67; 05: 0.33; 06: NA; 07: 1.03; 08: 0; 09: 0.97; 10: 0.06; 11: 0.76; 12: 1.51; 13: 0.44
| Analytical results for K in OREAS 600 (Certified Value 0.209 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  | 3
| 1 | 1 | 0.33 | 0.3333 | 0.21 | 0.2 | 0.2 | 0.2 | 0.31395 | 0.22 | 0.26 | 0.23531 | 0.203 | 0.25 | 0.2 | 0.24 | 0.16 | 0.177 | 0.2 | 0.19 | 0.21 | 0.19 |  |  | 1
| 1 | 2 | 0.33 | 0.3328 | 0.22 | 0.2 | 0.21 | 0.19 | 0.31962 | 0.22 | 0.25 | 0.23271 | 0.206 | 0.27 | 0.2 | 0.24 | 0.16 | 0.175 | 0.2 | 0.19 | 0.21 | 0.19 |  |  | 0
| 1 | 3 | 0.34 | 0.3337 | 0.21 | 0.2 | 0.2 | 0.2 | 0.32298 | 0.21 | 0.25 | 0.24018 | 0.204 | 0.27 | 0.2 | 0.24 | 0.16 | 0.178 | 0.2 | 0.18 | 0.2 | 0.19 |  |  | 16
| 1 | 4 | 0.33 | 0.3195 | 0.22 | 0.2 | 0.2 | 0.19 | 0.30239 | 0.22 | 0.27 | 0.24347 | 0.205 | 0.26 | 0.2 | 0.24 | 0.16 | 0.177 | 0.2 | 0.19 | 0.19 | 0.19 |  |  | 0.20886
| 1 | 5 | 0.34 | 0.3244 | 0.21 | 0.2 | 0.2 | 0.19 | 0.31473 | 0.21 | 0.26 | 0.24832 | 0.209 | 0.27 | 0.21 | 0.24 | 0.16 | 0.179 | 0.21 | 0.18 | 0.2 | 0.18 |  |  | 90
| 1 | 6 | 0.33 | 0.3315 | 0.21 | 0.19 | 0.21 | 0.19 | 0.30206 | 0.23 | 0.26 | 0.23873 | 0.198 | 0.27 | 0.21 | 0.24 | 0.16 | 0.174 | 0.21 | 0.18 | 0.2 | 0.18
| Mean |  | 0.33333 | 0.3292 | 0.21333 | 0.19833 | 0.20333 | 0.19333 | 0.31262 | 0.21833 | 0.25833 | 0.23979 | 0.20417 | 0.265 | 0.20333 | 0.24 | 0.16 | 0.17667 | 0.20333 | 0.185 | 0.20167 | 0.18667
| Median |  | 0.33 | 0.33215 | 0.21 | 0.2 | 0.2 | 0.19 | 0.31434 | 0.22 | 0.26 | 0.23946 | 0.2045 | 0.27 | 0.2 | 0.24 | 0.16 | 0.177 | 0.2 | 0.185 | 0.2 | 0.19
| Std Dev. |  | 0.005164 | 0.0058726 | 0.005164 | 0.0040825 | 0.005164 | 0.005164 | 0.0087016 | 0.0075277 | 0.0075277 | 0.0056182 | 0.003656 | 0.0083666 | 0.005164 | 0 | 0 | 0.0018619 | 0.005164 | 0.0054772 | 0.0075277 | 0.005164
| Rel.Std.Dev. |  | 1.549% | 1.784% | 2.421% | 2.058% | 2.54% | 2.671% | 2.783% | 3.448% | 2.914% | 2.343% | 1.791% | 3.157% | 2.54% | 0% | 0% | 1.054% | 2.54% | 2.961% | 3.733% | 2.766%
| PDM3 |  | 59.6% | 57.62% | 2.142% | -5.04% | -2.646% | -7.434% | 49.68% | 4.536% | 23.69% | 14.81% | -2.247% | 26.88% | -2.646% | 14.91% | -23.39% | -15.41% | -2.646% | -11.42% | -3.444% | -10.63%
| Z-Score (Absolute) |  | 4.88 | 4.72 | 0.36 | 0.2 | 0.02 | 0.39 | 4.1 | 0.55 | 2.05 | 1.36 | 0.02 | 2.31 | 0.02 | 1.36 | 1.65 | 1.02 | 0.02 | 0.71 | 0.08 | 0.64
| Analytical results for La in OREAS 600 (Certified Value 18.3 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 25 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g |  |  |  |  | 2
| 1 | 1 | 18.6912 | 17.95 | 10 | 20 | 20 | 20 | 19.4813 | 19 | 18 | 17.1883 | 19 | 18 | 17.1 | 16.8 | 15 | 17.48 | 19 | 18.8 |  |  |  |  | 1
| 1 | 2 | 18.5832 | 17.87 | 20 | 20 | 20 | 20 | 19.6927 | 18 | 17 | 17.6696 | 18 | 19 | 17 | 17.6 | 15 | 17.59 | 19 | 18.5 |  |  |  |  | 0
| 1 | 3 | 18.414 | 18.14 | 20 | 20 | 20 | 20 | 19.1911 | 17 | 17 | 17.5835 | 18 | 19 | 17 | 13.3 | 15 | 17.66 | 18 | 18.8 |  |  |  |  | 16
| 1 | 4 | 18.558 | 17.91 | 20 | 20 | 20 | 20 | 19.5217 | 18 | 17 | 18.3653 | 18 | 18 | 16.8 | 20.3 | 15 | 17.31 | 18 | 18.2 |  |  |  |  | 18.2764
| 1 | 5 | 18.6012 | 17.83 | 10 | 20 | 20 | 20 | 19.6784 | 17 | 17 | 17.9451 | 19 | 19 | 17.2 | 15.5 | 15 | 17.36 | 18 | 18.5 |  |  |  |  | 91
| 1 | 6 | 18.864 | 18.33 | 10 | 20 | 20 | 20 | 19.6051 | 18 | 17 | 18.2613 | 18 | 19 | 16.9 | 16.6 | 15 | 17.13 | 18 | 18.5
| Mean |  | 18.6186 | 18.005 | 15 | 20 | 20 | 20 | 19.5284 | 17.8333 | 17.1667 | 17.8355 | 18.3333 | 18.6667 | 17 | 16.6833 | 15 | 17.4217 | 18.3333 | 18.55
| Median |  | 18.5922 | 17.93 | 15 | 20 | 20 | 20 | 19.5634 | 18 | 17 | 17.8073 | 18 | 19 | 17 | 16.7 | 15 | 17.42 | 18 | 18.5
| Std Dev. |  | 0.15005 | 0.19222 | 5.4772 | 0 | 0 | 0 | 0.18515 | 0.75277 | 0.40825 | 0.4436 | 0.5164 | 0.5164 | 0.14142 | 2.3147 | 0 | 0.19488 | 0.5164 | 0.22583
| Rel.Std.Dev. |  | 0.8059% | 1.068% | 36.51% | 0% | 0% | 0% | 0.9481% | 4.221% | 2.378% | 2.487% | 2.817% | 2.766% | 0.8319% | 13.87% | 0% | 1.119% | 2.817% | 1.217%
| PDM3 |  | 1.872% | -1.485% | -17.93% | 9.431% | 9.431% | 9.431% | 6.85% | -2.424% | -6.072% | -2.412% | 0.3113% | 2.135% | -6.984% | -8.717% | -17.93% | -4.677% | 0.3113% | 1.497%
| Z-Score (Absolute) |  | 0.19 | 0.21 | 1.08 | 1.08 | 1.08 | 1.08 | 0.78 | 0.32 | 0.76 | 0.32 | 0 | 0.22 | 0.87 | 1.07 | 2.16 | 0.59 | 0 | 0.14
| Analytical results for Li in OREAS 600 (Certified Value 6.54 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [9 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 40g | [9 empty] | 2
| 1 | 1 | 8.459 | 7.1 | 8.6214 | 5 | 6.2476 | 6.6 | 6 | 6 | 7 | 6 | 6.632 | 5 | 6.2 | [9 empty] | 1
| 1 | 2 | 8.3215 | 7.1 | 8.4099 | 5 | 6.3355 | 7 | 6 | 6 | 7 | 6 | 6.764 | 5 | 5.9 | [9 empty] | 0
| 1 | 3 | 8.316 | 6.8 | 8.3023 | 5 | 6.1987 | 7.3 | 6 | 6 | 7 | 6 | 7.193 | 5 | 5.8 | [9 empty] | 16
| 1 | 4 | 8.459 | 7 | 8.3183 | 5 | 6.434 | 6.7 | 6 | 6 | 7 | 6 | 7.343 | 5 | 5.5 | [9 empty] | 6.5375
| 1 | 5 | 8.866 | 6.8 | 8.6973 | 5 | 6.3591 | 7.1 | 6 | 6 | 7 | 6 | 7.803 | 6 | 5.8 | [9 empty] | 92
| 1 | 6 | 8.4645 | 7 | 8.4389 | 5 | 6.1846 | 6.8 | 6 | 6 | 7 | 5 | 6.522 | 6 | 5.2
| Mean |  | 8.481 | 6.9667 | 8.4647 | 5 | 6.2932 | 6.9167 | 6 | 6 | 7 | 5.8333 | 7.0428 | 5.3333 | 5.7333
| Median |  | 8.459 | 7 | 8.4244 | 5 | 6.2915 | 6.9 | 6 | 6 | 7 | 6 | 6.9785 | 5 | 5.8
| Std Dev. |  | 0.20106 | 0.13663 | 0.16131 | 0 | 0.098747 | 0.26394 | 0 | 0 | 0 | 0.40825 | 0.49159 | 0.5164 | 0.34448
| Rel.Std.Dev. |  | 2.371% | 1.961% | 1.906% | 0% | 1.569% | 3.816% | 0% | 0% | 0% | 6.999% | 6.98% | 9.682% | 6.008%
| PDM3 |  | 29.73% | 6.564% | 29.48% | -23.52% | -3.737% | 5.799% | -8.222% | -8.222% | 7.074% | -10.77% | 7.729% | -18.42% | -12.3%
| Z-Score (Absolute) |  | 2.11 | 0.67 | 2.17 | 1.29 | 0 | 0.62 | 0.29 | 0.29 | 0.71 | 0.46 | 0.75 | 0.96 | 0.56
| Analytical results for Lu in OREAS 600 (Certified Value 0.063 ppm) | [23 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 14; 06: 15; 07: 17; 08: 19; 09: 25 |  |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*MS |  |  |  |  | 3
|  |  | 01: 40g; 02: 20g; 05: 0.5g; 06: 01g; 07: 0.25g; 08: 0.25g; 09: 40g |  |  |  |  | 3
| 1 | 1 | 01: 0.06; 02: 0.057648; 05: 0.065; 06: < 0.1; 07: 0.09; 08: 0.064; 09: 0.06 |  |  |  |  | 1
| 1 | 2 | 01: 0.06; 02: 0.060478; 05: 0.061; 06: < 0.1; 07: 0.1; 08: 0.064; 09: 0.05 |  |  |  |  | 0
| 1 | 3 | 01: 0.06; 02: 0.055099; 05: 0.064; 06: < 0.1; 07: 0.07; 08: 0.065; 09: 0.06 |  |  |  |  | 16
| 1 | 4 | 01: 0.07; 02: 0.060199; 05: 0.06; 06: < 0.1; 07: 0.07; 08: 0.065; 09: 0.06 |  |  |  |  | 0.062557
| 1 | 5 | 01: 0.06; 02: 0.05702; 05: 0.062; 06: < 0.1; 07: 0.07; 08: 0.071; 09: 0.06 |  |  |  |  | 93
| 1 | 6 | 01: 0.06; 02: 0.05561; 05: 0.062; 06: < 0.1; 07: 0.07; 08: 0.063; 09: 0.06
| Mean |  | 01: 0.061667; 02: 0.057676; 05: 0.062333; 07: 0.078333; 08: 0.065333; 09: 0.058333
| Median |  | 01: 0.06; 02: 0.057334; 05: 0.062; 07: 0.07; 08: 0.0645; 09: 0.06
| Std Dev. |  | 01: 0.0040825; 02: 0.0022609; 05: 0.0018619; 07: 0.013292; 08: 0.0028752; 09: 0.0040825
| Rel.Std.Dev. |  | 01: 6.62%; 02: 3.92%; 05: 2.987%; 07: 16.97%; 08: 4.401%; 09: 6.999%
| PDM3 |  | 01: -1.423%; 02: -7.803%; 05: -0.3576%; 07: 25.22%; 08: 4.438%; 09: -6.752%
| Z-Score (Absolute) |  | 01: 0.06; 02: 0.83; 05: 0.06; 06: NA; 07: 1.54; 08: 0.64; 09: 0.71
| Analytical results for Mg in OREAS 600 (Certified Value 0.342 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 0.46667 | 0.42 | 0.32 | 0.32 | 0.32 | 0.31 | 0.43455 | 0.225 | 0.3 | 0.36 | 0.38811 | 0.33 | 0.4 | 0.31 | 0.37 | 0.26 | 0.365 | 0.26 | 0.32 | 0.39 | 0.32 |  | 1
| 1 | 2 | 0.45667 | 0.42 | 0.33 | 0.32 | 0.32 | 0.3 | 0.4273 | 0.25 | 0.29 | 0.35 | 0.38453 | 0.32 | 0.38 | 0.32 | 0.37 | 0.26 | 0.375 | 0.26 | 0.32 | 0.38 | 0.32 |  | 0
| 1 | 3 | 0.45667 | 0.43 | 0.32 | 0.32 | 0.32 | 0.31 | 0.42468 | 0.25 | 0.3 | 0.35 | 0.39541 | 0.33 | 0.38 | 0.31 | 0.33 | 0.27 | 0.374 | 0.26 | 0.31 | 0.39 | 0.32 |  | 16
| 1 | 4 | 0.46 | 0.42 | 0.33 | 0.32 | 0.32 | 0.31 | 0.43152 | 0.215 | 0.29 | 0.35 | 0.39567 | 0.33 | 0.39 | 0.32 | 0.37 | 0.26 | 0.372 | 0.27 | 0.32 | 0.4 | 0.31 |  | 0.3421
| 1 | 5 | 0.46667 | 0.42 | 0.32 | 0.32 | 0.31 | 0.3 | 0.44057 | 0.235 | 0.29 | 0.35 | 0.39936 | 0.34 | 0.39 | 0.32 | 0.37 | 0.26 | 0.377 | 0.28 | 0.31 | 0.37 | 0.31 |  | 94
| 1 | 6 | 0.46 | 0.42 | 0.33 | 0.31 | 0.32 | 0.3 | 0.42903 | 0.255 | 0.29 | 0.35 | 0.39492 | 0.33 | 0.39 | 0.32 | 0.37 | 0.27 | 0.369 | 0.28 | 0.31 | 0.41 | 0.3
| Mean |  | 0.46111 | 0.42167 | 0.325 | 0.31833 | 0.31833 | 0.305 | 0.43127 | 0.23833 | 0.29333 | 0.35167 | 0.393 | 0.33 | 0.38833 | 0.31667 | 0.36333 | 0.26333 | 0.372 | 0.26833 | 0.315 | 0.39 | 0.31333
| Median |  | 0.46 | 0.42 | 0.325 | 0.32 | 0.32 | 0.305 | 0.43027 | 0.2425 | 0.29 | 0.35 | 0.39517 | 0.33 | 0.39 | 0.32 | 0.37 | 0.26 | 0.373 | 0.265 | 0.315 | 0.39 | 0.315
| Std Dev. |  | 0.0045542 | 0.0040825 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0056819 | 0.016021 | 0.005164 | 0.0040825 | 0.0055273 | 0.0063246 | 0.0075277 | 0.005164 | 0.01633 | 0.005164 | 0.0043818 | 0.0098319 | 0.0054772 | 0.014142 | 0.008165
| Rel.Std.Dev. |  | 0.9877% | 0.9682% | 1.685% | 1.282% | 1.282% | 1.796% | 1.317% | 6.722% | 1.76% | 1.161% | 1.406% | 1.917% | 1.938% | 1.631% | 4.494% | 1.961% | 1.178% | 3.664% | 1.739% | 3.626% | 2.606%
| PDM3 |  | 34.79% | 23.26% | -4.998% | -6.947% | -6.947% | -10.84% | 26.07% | -30.33% | -14.25% | 2.797% | 14.88% | -3.536% | 13.52% | -7.434% | 6.207% | -23.02% | 8.741% | -21.56% | -7.921% | 14% | -8.408%
| Z-Score (Absolute) |  | 2.04 | 1.45 | 0 | 0.1 | 0.1 | 0.3 | 1.59 | 1.3 | 0.47 | 0.4 | 1.02 | 0.07 | 0.95 | 0.12 | 0.67 | 0.92 | 0.7 | 0.85 | 0.15 | 0.97 | 0.17
| Analytical results for Mn in OREAS 600 (Certified Value 0.068 wt.%) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 0.070857 | 0.068 | 0.0662 | 0.0686 | 0.0691 | 0.0677 | 0.074016 | 0.072 | 0.0624 | 0.0712 | 0.066733 | 0.0701 | 0.07224 | 0.0641 | 0.0683 | 0.06 | 0.070398 | 0.0642 | 0.0668 | 0.072 | 0.0706 |  | 1
| 1 | 2 | 0.071655 | 0.0698 | 0.0686 | 0.0683 | 0.0668 | 0.0681 | 0.072632 | 0.0746 | 0.0617 | 0.0685 | 0.06689 | 0.0684 | 0.072975 | 0.0648 | 0.0685 | 0.06 | 0.070173 | 0.0653 | 0.0662 | 0.07 | 0.0709 |  | 0
| 1 | 3 | 0.071632 | 0.0704 | 0.0673 | 0.0684 | 0.0666 | 0.0662 | 0.07216 | 0.0745 | 0.0629 | 0.0681 | 0.067181 | 0.0702 | 0.07455 | 0.0641 | 0.067 | 0.06 | 0.071894 | 0.0645 | 0.0656 | 0.07 | 0.0709 |  | 16
| 1 | 4 | 0.070962 | 0.0702 | 0.0694 | 0.0689 | 0.0693 | 0.0658 | 0.07322 | 0.066 | 0.0625 | 0.0697 | 0.067953 | 0.0697 | 0.072765 | 0.0646 | 0.0685 | 0.06 | 0.070688 | 0.0662 | 0.0664 | 0.07 | 0.07 |  | 0.068354
| 1 | 5 | 0.071489 | 0.0693 | 0.0655 | 0.0688 | 0.0665 | 0.0679 | 0.074617 | 0.0723 | 0.0603 | 0.0696 | 0.067574 | 0.0715 | 0.073185 | 0.0647 | 0.0656 | 0.06 | 0.071941 | 0.0673 | 0.0655 | 0.067 | 0.0708 |  | 95
| 1 | 6 | 0.070916 | 0.0698 | 0.0682 | 0.0676 | 0.0685 | 0.0658 | 0.073998 | 0.0725 | 0.0626 | 0.0693 | 0.066724 | 0.0689 | 0.0735 | 0.0638 | 0.0655 | 0.06 | 0.070684 | 0.0683 | 0.0648 | 0.073 | 0.0702
| Mean |  | 0.071252 | 0.069583 | 0.067533 | 0.068433 | 0.0678 | 0.066917 | 0.07344 | 0.071983 | 0.062067 | 0.0694 | 0.067176 | 0.0698 | 0.073203 | 0.06435 | 0.067233 | 0.06 | 0.070963 | 0.065967 | 0.065883 | 0.070333 | 0.070567
| Median |  | 0.071225 | 0.0698 | 0.06775 | 0.0685 | 0.06765 | 0.06695 | 0.073609 | 0.0724 | 0.06245 | 0.06945 | 0.067035 | 0.0699 | 0.07308 | 0.06435 | 0.06765 | 0.06 | 0.070686 | 0.06575 | 0.0659 | 0.07 | 0.0707
| Std Dev. |  | 0.00037825 | 0.00086352 | 0.0014855 | 0.00046762 | 0.0013084 | 0.0010944 | 0.00093456 | 0.0031416 | 0.00095219 | 0.0010844 | 0.00049882 | 0.0010881 | 0.00078411 | 0.00040373 | 0.0014194 | 0 | 0.0007642 | 0.001612 | 0.00072226 | 0.0020656 | 0.00038297
| Rel.Std.Dev. |  | 0.5309% | 1.241% | 2.2% | 0.6833% | 1.93% | 1.635% | 1.273% | 4.364% | 1.534% | 1.563% | 0.7426% | 1.559% | 1.071% | 0.6274% | 2.111% | 0% | 1.077% | 2.444% | 1.096% | 2.937% | 0.5427%
| PDM3 |  | 4.24% | 1.799% | -1.201% | 0.1161% | -0.8105% | -2.103% | 7.441% | 5.31% | -9.198% | 1.53% | -1.724% | 2.115% | 7.093% | -5.858% | -1.639% | -12.22% | 3.817% | -3.493% | -3.614% | 2.896% | 3.237%
| Z-Score (Absolute) |  | 0.89 | 0.36 | 0.28 | 0 | 0.2 | 0.48 | 1.58 | 1.5 | 1.9 | 0.31 | 0.4 | 0.43 | 1.51 | 1.29 | 0.38 | 2.67 | 0.8 | 0.78 | 0.81 | 0.6 | 0.67
| Analytical results for Mo in OREAS 600 (Certified Value 1.92 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 2
| 1 | 1 | 2.3433 | 2 | 2 | 3 | 2 | 2 | 2.077 | 2 | 2 | 2 | 1.3183 | 2 | 2 | 3 | 2.12 | 3 | 1.999 | 2 | 2 | 1.8 | 2.4 |  | 1
| 1 | 2 | 2.0333 | 2.5 | 2 | 4 | 2 | 2 | 1.8485 | 2 | 2 | 2 | 1.3146 | 2 | 2 | 6 | 1.48 | 2 | 2.013 | 2 | 2 | 2 | 2.2 |  | 0
| 1 | 3 | 1.8433 | 2.4 | 2 | 3 | 2 | 3 | 1.8353 | 2 | 2 | 1 | 1.3317 | 2 | 2 | 3 | 1.17 | 3 | 1.877 | 2 | 1 | 1.8 | 2.2 |  | 16
| 1 | 4 | 1.91 | 2.6 | 2 | 4 | 2 | 2 | 1.8609 | 2 | 2 | 2 | 1.3464 | 2 | 2 | 4 | 1.28 | 3 | 1.889 | 1 | 2 | 1.8 | 2.4 |  | 1.9153
| 1 | 5 | 1.85 | 2.5 | 2 | 3 | 2 | 2 | 1.9118 | 2 | 2 | 2 | 1.4746 | 2 | 2 | 3 | 1.13 | 4 | 1.969 | 2 | 2 | 2 | 2.4 |  | 96
| 1 | 6 | 1.86 | 2.4 | 2 | 3 | 2 | 2 | 1.9315 | 2 | 2 | 2 | 1.3201 | 2 | 2 | 3 | 1.16 | 3 | 1.993 | 2 | 2 | 9.2 | 2.4
| Mean |  | 1.9733 | 2.4 | 2 | 3.3333 | 2 | 2.1667 | 1.9108 | 2 | 2 | 1.8333 | 1.351 | 2 | 2 | 3.6667 | 1.39 | 3 | 1.9567 | 1.8333 | 1.8333 | 3.1 | 2.3333
| Median |  | 1.885 | 2.45 | 2 | 3 | 2 | 2 | 1.8863 | 2 | 2 | 2 | 1.3259 | 2 | 2 | 3 | 1.225 | 3 | 1.981 | 2 | 2 | 1.9 | 2.4
| Std Dev. |  | 0.19466 | 0.20976 | 0 | 0.5164 | 0 | 0.40825 | 0.089604 | 0 | 0 | 0.40825 | 0.061663 | 0 | 0 | 1.2111 | 0.38 | 0.63246 | 0.058929 | 0.40825 | 0.40825 | 2.99 | 0.10328
| Rel.Std.Dev. |  | 9.865% | 8.74% | 0% | 15.49% | 0% | 18.84% | 4.689% | 0% | 0% | 22.27% | 4.564% | 0% | 0% | 33.03% | 27.34% | 21.08% | 3.012% | 22.27% | 22.27% | 96.45% | 4.426%
| PDM3 |  | 3.031% | 25.31% | 4.423% | 74.04% | 4.423% | 13.13% | -0.2329% | 4.423% | 4.423% | -4.279% | -29.46% | 4.423% | 4.423% | 91.44% | -27.43% | 56.63% | 2.161% | -4.279% | -4.279% | 61.86% | 21.83%
| Z-Score (Absolute) |  | 0.57 | 2.25 | 0 | 7.49 | 0 | 0 | 0.5 | 0 | 0 | 0.94 | 3.65 | 0 | 0 | 6.74 | 4.25 | 5.62 | 0.24 | 0.94 | 0.94 | 0.67 | 1.87
| Analytical results for Na in OREAS 600 (Certified Value 0.051 wt.%) | [23 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  |  | 3
| 1 | 1 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.050245 | 0.05 | 0.07 | 0.049734 | 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | 0.061 | 0.08 | 0.04 | 0.05 |  |  |  | 1
| 1 | 2 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.049344 | 0.05 | 0.06 | 0.049876 | 0.05 | 0.06 | 0.05 | 0.06 | 0.05 | 0.058 | 0.08 | 0.04 | 0.05 |  |  |  | 0
| 1 | 3 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.048326 | 0.05 | 0.06 | 0.050481 | 0.05 | 0.06 | 0.05 | 0.05 | 0.05 | 0.062 | 0.08 | 0.04 | 0.05 |  |  |  | 16
| 1 | 4 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.050433 | 0.05 | 0.07 | 0.05077 | 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | 0.06 | 0.08 | 0.04 | 0.05 |  |  |  | 0.050627
| 1 | 5 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.051018 | 0.05 | 0.06 | 0.050847 | 0.05 | 0.06 | 0.05 | 0.06 | 0.05 | 0.062 | 0.08 | 0.04 | 0.05 |  |  |  | 97
| 1 | 6 | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.048674 | 0.05 | 0.06 | 0.050182 | 0.05 | 0.06 | 0.05 | 0.06 | 0.05 | 0.061 | 0.08 | 0.04 | 0.05
| Mean |  | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.049673 | 0.05 | 0.063333 | 0.050315 | 0.05 | 0.056667 | 0.05 | 0.058333 | 0.05 | 0.060667 | 0.08 | 0.04 | 0.05
| Median |  | 0.04 | 0.05 | 0.03 | 0.05 | 0.05 | 0.04 | 0.049795 | 0.05 | 0.06 | 0.050332 | 0.05 | 0.06 | 0.05 | 0.06 | 0.05 | 0.061 | 0.08 | 0.04 | 0.05
| Std Dev. |  | 0 | 7.6012e-18 | 0 | 7.6012e-18 | 7.6012e-18 | 0 | 0.0010615 | 7.6012e-18 | 0.005164 | 0.00046176 | 7.6012e-18 | 0.005164 | 7.6012e-18 | 0.0040825 | 7.6012e-18 | 0.0015055 | 0 | 0 | 7.6012e-18
| Rel.Std.Dev. |  | 0% | 1.52e-14% | 0% | 1.52e-14% | 1.52e-14% | 0% | 2.137% | 1.52e-14% | 8.154% | 0.9177% | 1.52e-14% | 9.113% | 1.52e-14% | 6.999% | 1.52e-14% | 2.482% | 0% | 0% | 1.52e-14%
| PDM3 |  | -20.99% | -1.238% | -40.74% | -1.238% | -1.238% | -20.99% | -1.883% | -1.238% | 25.1% | -0.6157% | -1.238% | 11.93% | -1.238% | 15.22% | -1.238% | 19.83% | 58.02% | -20.99% | -1.238%
| Z-Score (Absolute) |  | 20.64 | 0 | 41.29 | 0 | 0 | 20.64 | 0.67 | 0 | 27.53 | 0.65 | 0 | 13.76 | 0 | 20.64 | 0 | 22.02 | 61.93 | 20.64 | 0
| Analytical results for Nb in OREAS 600 (Indicative Value 0.18 ppm) | [23 empty] | Indicative
Nb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 14; 08: 15; 09: 17; 11: 19; 12: 25 |  |  |  | ppm
|  |  | 01: AR*OES; 02: AR*MS; 05: AR*OES/MS; 06: AR*OES; 07: AR*MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*OES/MS; 12: AR*MS |  |  |  | 2
|  |  | 01: 0.5g; 02: 40g; 05: 20g; 06: 0.15g; 07: 0.5g; 08: 01g; 09: 0.25g; 11: 0.25g; 12: 40g |  |  |  | 2
| 1 | 1 | 01: < 5; 02: 0.35; 05: 0.06759; 06: < 10; 07: 0.08; 08: < 0.2; 09: 0.22; 11: 0.236; 12: 0.2 |  |  |  | 1
| 1 | 2 | 01: < 5; 02: 0.4; 05: 0.075557; 06: < 10; 07: 0.07; 08: < 0.2; 09: 0.16; 11: 0.24; 12: 0.2 |  |  |  | 8
| 1 | 3 | 01: < 5; 02: 0.35; 05: 0.057802; 06: < 10; 07: 0.1; 08: < 0.2; 09: 0.13; 11: 0.253; 12: 0.2 |  |  |  | 16
| 1 | 4 | 01: < 5; 02: 0.4; 05: 0.062571; 06: < 10; 07: 0.06; 08: < 0.2; 09: 0.14; 11: 0.242; 12: 0.2 |  |  |  | 0.17942
| 1 | 5 | 01: < 5; 02: 0.34; 05: 0.074265; 06: < 10; 07: 0.06; 08: < 0.2; 09: 0.11; 11: 0.235; 12: 0.2 |  |  |  | 14
| 1 | 6 | 01: < 5; 02: 0.32; 05: 0.08219; 06: < 10; 07: 0.07; 08: < 0.2; 09: 0.13; 11: 0.229; 12: < 0.2
| Mean |  | 02: 0.36; 05: 0.069996; 07: 0.073333; 09: 0.14833; 11: 0.23917; 12: 0.2
| Median |  | 02: 0.35; 05: 0.070927; 07: 0.07; 09: 0.135; 11: 0.238; 12: 0.2
| Std Dev. |  | 02: 0.032863; 05: 0.0090275; 07: 0.015055; 09: 0.038687; 11: 0.0081343; 12: 0
| Rel.Std.Dev. |  | 02: 9.129%; 05: 12.9%; 07: 20.53%; 09: 26.08%; 11: 3.401%; 12: 0%
| PDM3 |  | 02: 100.7%; 05: -60.99%; 07: -59.13%; 09: -17.32%; 11: 33.3%; 12: 11.47%
| Z-Score (Absolute) |  | 01: NA; 02: 1.57; 05: 0.79; 06: NA; 07: 0.76; 08: NA; 09: 0.27; 11: 0.59; 12: 0.27
| Analytical results for Nd in OREAS 600 (Certified Value 14.1 ppm) | [23 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  |  |  | 1
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  |  |  | 2
| 1 | 1 | 01: 13.61; 02: 15.7373; 05: 12.3511; 06: 14.18; 07: 15.3; 08: 13.7 |  |  |  | 1
| 1 | 2 | 01: 13.61; 02: 15.5279; 05: 12.0547; 06: 14.05; 07: 14.43; 08: 13.9 |  |  |  | 0
| 1 | 3 | 01: 14.01; 02: 15.6639; 05: 12.4917; 06: 14.21; 07: 14.82; 08: 13.8 |  |  |  | 16
| 1 | 4 | 01: 13.9; 02: 15.6567; 05: 12.3927; 06: 13.8; 07: 14.85; 08: 13.5 |  |  |  | 14.0837
| 1 | 5 | 01: 13.72; 02: 15.7029; 05: 12.6343; 06: 14.06; 07: 14.76; 08: 13.7 |  |  |  | 98
| 1 | 6 | 01: 14.08; 02: 15.8198; 05: 12.6997; 06: 13.6; 07: 15.19; 08: 13.5
| Mean |  | 01: 13.8217; 02: 15.6848; 05: 12.4374; 06: 13.9833; 07: 14.8917; 08: 13.6833
| Median |  | 01: 13.81; 02: 15.6834; 05: 12.4422; 06: 14.055; 07: 14.835; 08: 13.7
| Std Dev. |  | 01: 0.2041; 02: 0.097094; 05: 0.23083; 06: 0.23704; 07: 0.31403; 08: 0.16021
| Rel.Std.Dev. |  | 01: 1.477%; 02: 0.619%; 05: 1.856%; 06: 1.695%; 07: 2.109%; 08: 1.171%
| PDM3 |  | 01: -1.86%; 02: 11.37%; 05: -11.69%; 06: -0.7126%; 07: 5.737%; 08: -2.843%
| Z-Score (Absolute) |  | 01: 0.09; 02: 1.99; 05: 1.64; 06: 0.09; 07: 1.1; 08: 0.24
| Analytical results for Ni in OREAS 600 (Certified Value 15.4 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 1
| 1 | 1 | 19.125 | 16.2 | 15 | 18 | 16 | 18 | 16.3807 | 24 | 17 | 15 | 13.0677 | 15 | 15 | 14 | 12.6 | 14 | 15.776 | 14 | 15 | 15 | 15 |  | 1
| 1 | 2 | 18.18 | 17.8 | 15 | 16 | 19 | 17 | 15.8421 | 42 | 20 | 15 | 13.8219 | 15 | 14 | 16 | 14.2 | 14 | 14.953 | 13 | 14 | 17 | 15 |  | 0
| 1 | 3 | 17.78 | 15.7 | 14 | 14 | 18 | 17 | 15.0684 | 23 | 16 | 16 | 12.8837 | 18 | 14 | 13 | 10.8 | 13 | 13.879 | 15 | 17 | 16 | 15 |  | 16
| 1 | 4 | 17.365 | 17.9 | 15 | 27 | 18 | 14 | 17.2601 | 21 | 17 | 15 | 13.5091 | 16 | 16 | 15 | 12.8 | 12 | 13.51 | 15 | 15 | 17 | 14 |  | 15.354
| 1 | 5 | 19.39 | 17.2 | 14 | 16 | 17 | 15 | 16.1441 | 24 | 16 | 15 | 12.6545 | 15 | 14 | 13 | 12.6 | 14 | 15.161 | 15 | 15 | 17 | 14 |  | 99
| 1 | 6 | 17.27 | 15.9 | 15 | 16 | 20 | 16 | 17.1646 | 25 | 17 | 17 | 13.825 | 13 | 14 | 12 | 15.5 | 12 | 14.666 | 14 | 17 | 17 | 15
| Mean |  | 18.185 | 16.7833 | 14.6667 | 17.8333 | 18 | 16.1667 | 16.31 | 26.5 | 17.1667 | 15.5 | 13.2936 | 15.3333 | 14.5 | 13.8333 | 13.0833 | 13.1667 | 14.6575 | 14.3333 | 15.5 | 16.5 | 14.6667
| Median |  | 17.98 | 16.7 | 15 | 16 | 18 | 16.5 | 16.2624 | 24 | 17 | 15 | 13.2884 | 15 | 14 | 13.5 | 12.7 | 13.5 | 14.8095 | 14.5 | 15 | 17 | 15
| Std Dev. |  | 0.89547 | 0.97451 | 0.5164 | 4.6655 | 1.4142 | 1.472 | 0.82776 | 7.7136 | 1.472 | 0.83666 | 0.49708 | 1.633 | 0.83666 | 1.472 | 1.603 | 0.98319 | 0.83842 | 0.8165 | 1.2247 | 0.83666 | 0.5164
| Rel.Std.Dev. |  | 4.924% | 5.806% | 3.521% | 26.16% | 7.857% | 9.105% | 5.075% | 29.11% | 8.575% | 5.398% | 3.739% | 10.65% | 5.77% | 10.64% | 12.25% | 7.467% | 5.72% | 5.696% | 7.902% | 5.071% | 3.521%
| PDM3 |  | 18.44% | 9.309% | -4.476% | 16.15% | 17.23% | 5.293% | 6.227% | 72.59% | 11.81% | 0.9511% | -13.42% | -0.1344% | -5.562% | -9.904% | -14.79% | -14.25% | -4.536% | -6.647% | 0.9511% | 7.464% | -4.476%
| Z-Score (Absolute) |  | 1.81 | 0.87 | 0.56 | 0.34 | 1.69 | 0.45 | 0.55 | 5.33 | 0.74 | 0 | 1.49 | 0.11 | 0.67 | 1.12 | 1.63 | 1.57 | 0.57 | 0.79 | 0 | 0.67 | 0.56
| Analytical results for P in OREAS 600 (Certified Value 0.05 wt.%) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  |  | 3
| 1 | 1 | 0.053018 | 0.048 | 0.048 | 0.051 | 0.051 | 0.05 | 0.051425 | 0.0485 | 0.0574 | 0.047313 | 0.0502 | 0.05 | 0.054 | 0.03 | 0.054109 | 0.0471 | 0.048 | 0.048 | 0.048 |  |  |  | 1
| 1 | 2 | 0.053587 | 0.0475 | 0.05 | 0.051 | 0.05 | 0.048 | 0.053412 | 0.049 | 0.0558 | 0.047639 | 0.0501 | 0.04 | 0.055 | 0.03 | 0.047999 | 0.0475 | 0.048 | 0.054 | 0.048 |  |  |  | 0
| 1 | 3 | 0.053157 | 0.0487 | 0.049 | 0.051 | 0.051 | 0.05 | 0.048937 | 0.0489 | 0.0559 | 0.047231 | 0.0499 | 0.03 | 0.05 | 0.03 | 0.062054 | 0.0461 | 0.047 | 0.056 | 0.05 |  |  |  | 16
| 1 | 4 | 0.052392 | 0.0482 | 0.05 | 0.051 | 0.051 | 0.049 | 0.049761 | 0.0479 | 0.0571 | 0.047981 | 0.0492 | 0.04 | 0.053 | 0.04 | 0.049102 | 0.0481 | 0.048 | 0.056 | 0.05 |  |  |  | 0.049827
| 1 | 5 | 0.053609 | 0.0458 | 0.048 | 0.051 | 0.049 | 0.05 | 0.051168 | 0.0481 | 0.0565 | 0.047494 | 0.0504 | 0.04 | 0.053 | 0.03 | 0.051882 | 0.0493 | 0.047 | 0.05 | 0.048 |  |  |  | 100
| 1 | 6 | 0.053091 | 0.0484 | 0.05 | 0.051 | 0.05 | 0.048 | 0.051697 | 0.0483 | 0.0566 | 0.047727 | 0.0489 | 0.06 | 0.053 | 0.04 | 0.064025 | 0.0491 | 0.046 | 0.054 | 0.05
| Mean |  | 0.053142 | 0.047767 | 0.049167 | 0.051 | 0.050333 | 0.049167 | 0.051067 | 0.04845 | 0.05655 | 0.047564 | 0.049783 | 0.043333 | 0.053 | 0.033333 | 0.054862 | 0.047867 | 0.047333 | 0.053 | 0.049
| Median |  | 0.053124 | 0.0481 | 0.0495 | 0.051 | 0.0505 | 0.0495 | 0.051296 | 0.0484 | 0.05655 | 0.047566 | 0.05 | 0.04 | 0.053 | 0.03 | 0.052995 | 0.0478 | 0.0475 | 0.054 | 0.049
| Std Dev. |  | 0.00044668 | 0.0010443 | 0.00098319 | 7.6012e-18 | 0.0008165 | 0.00098319 | 0.0015669 | 0.00043704 | 0.00063482 | 0.00027756 | 0.00059805 | 0.010328 | 0.0016733 | 0.005164 | 0.0067143 | 0.0012226 | 0.0008165 | 0.0032863 | 0.0010954
| Rel.Std.Dev. |  | 0.8405% | 2.186% | 2% | 1.49e-14% | 1.622% | 2% | 3.068% | 0.902% | 1.123% | 0.5836% | 1.201% | 23.83% | 3.157% | 15.49% | 12.24% | 2.554% | 1.725% | 6.201% | 2.236%
| PDM3 |  | 6.653% | -4.135% | -1.326% | 2.354% | 1.016% | -1.326% | 2.487% | -2.764% | 13.49% | -4.542% | -0.08801% | -13.03% | 6.368% | -33.1% | 10.1% | -3.935% | -5.005% | 6.368% | -1.66%
| Z-Score (Absolute) |  | 1.46 | 0.37 | 0 | 0.67 | 0.43 | 0 | 0.7 | 0.26 | 2.72 | 0.59 | 0.23 | 1.16 | 1.41 | 5.82 | 1.42 | 0.48 | 0.67 | 1.41 | 0.06
| Analytical results for Pb in OREAS 600 (Certified Value 157 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 0
| 1 | 1 | 156.41 | 157.9 | 157 | 157 | 160 | 156 | 149.0723 | 190 | 163 | 160 | 165.38 | 153 | 153 | 146 | 169 | 137 | 157.45 | 150 | 160 | 137 | 154 |  | 1
| 1 | 2 | 156.35 | 153 | 165 | 160 | 154 | 156 | 148.1014 | 190 | 160 | 156 | 169.2623 | 155 | 155 | 145 | 170 | 136 | 164.51 | 150 | 153 | 142 | 156 |  | 0
| 1 | 3 | 156.78 | 154.4 | 165 | 159 | 160 | 166 | 153.2669 | 180 | 156 | 157 | 171.1965 | 155 | 159 | 144 | 142 | 135 | 162.33 | 148 | 153 | 137 | 156 |  | 16
| 1 | 4 | 156.825 | 152.9 | 166 | 161 | 160 | 157 | 151.6702 | 160 | 156 | 161 | 171.3514 | 155 | 157 | 148 | 171 | 142 | 166 | 155 | 157 | 142 | 153 |  | 156.6304
| 1 | 5 | 158.36 | 155.1 | 156 | 157 | 158 | 151 | 156.5557 | 180 | 154 | 154 | 167.0009 | 154 | 160 | 146 | 154 | 135 | 154.54 | 158 | 153 | 144 | 156 |  | 101
| 1 | 6 | 161.845 | 153.9 | 157 | 153 | 163 | 151 | 150.2907 | 180 | 155 | 157 | 167.2376 | 154 | 156 | 159 | 163 | 146 | 150.26 | 149 | 156 | 140 | 159
| Mean |  | 157.7617 | 154.5333 | 161 | 157.8333 | 159.1667 | 156.1667 | 151.4929 | 180 | 157.3333 | 157.5 | 168.5714 | 154.3333 | 156.6667 | 148 | 161.5 | 138.5 | 159.1817 | 151.6667 | 155.3333 | 140.3333 | 155.6667
| Median |  | 156.8025 | 154.15 | 161 | 158 | 160 | 156 | 150.9804 | 180 | 156 | 157 | 168.2499 | 154.5 | 156.5 | 146 | 166 | 136.5 | 159.89 | 150 | 154.5 | 141 | 156
| Std Dev. |  | 2.1304 | 1.849 | 4.7749 | 2.8577 | 2.9944 | 5.4924 | 3.086 | 10.9545 | 3.4448 | 2.5884 | 2.4301 | 0.8165 | 2.582 | 5.5498 | 11.4673 | 4.5056 | 6.1471 | 3.9328 | 2.8752 | 2.8752 | 2.0656
| Rel.Std.Dev. |  | 1.35% | 1.196% | 2.966% | 1.811% | 1.881% | 3.517% | 2.037% | 6.086% | 2.189% | 1.643% | 1.442% | 0.529% | 1.648% | 3.75% | 7.101% | 3.253% | 3.862% | 2.593% | 1.851% | 2.049% | 1.327%
| PDM3 |  | 0.7223% | -1.339% | 2.79% | 0.768% | 1.619% | -0.2961% | -3.28% | 14.92% | 0.4488% | 0.5552% | 7.624% | -1.467% | 0.02316% | -5.51% | 3.109% | -11.58% | 1.629% | -3.169% | -0.8281% | -10.4% | -0.6153%
| Z-Score (Absolute) |  | 0.07 | 0.57 | 1.16 | 0.31 | 0.95 | 0.13 | 1.39 | 6.26 | 0.18 | 0.22 | 3.19 | 0.63 | 0 | 2.91 | 1.3 | 4.87 | 0.67 | 1.68 | 0.36 | 4.38 | 0.27
| Analytical results for Pd in OREAS 600 (Indicative Value < 10 ppb) | [23 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab |  | 1
| No | No | 01: 09; 02: 14; 05: 25 |  | ppb
|  |  | 01: AR*OES/MS; 02: AR*MS; 05: AR*MS |  | 0
|  |  | 01: 20g; 02: 0.5g; 05: 40g |  | 0
| 1 | 1 | 01: < 10; 02: < 5; 05: 10 |  | 1
| 1 | 2 | 01: < 10; 02: < 5; 05: < 10 |  | 9
| 1 | 3 | 01: < 10; 02: < 5; 05: < 10 |  | 16
| 1 | 4 | 01: < 10; 02: < 5; 05: < 10 |  | < 10
| 1 | 5 | 01: < 10; 02: < 5; 05: 10 |  | 15
| 1 | 6 | 01: < 10; 02: < 5; 05: 10
| Mean |  | 05: 10
| Median |  | 05: 10
| Std Dev. |  | 05: 0
Rel.Std.Dev. | bar: 05
PDM3
Z-Score (Absolute) | bar: 01 to 05
| Analytical results for Pr in OREAS 600 (Certified Value 4.04 ppm) | [23 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  | 3
| 1 | 1 | 01: 3.91; 02: 4.0726; 05: 3.949; 06: 3.83; 07: 4.29; 08: 3.62 |  | 1
| 1 | 2 | 01: 3.92; 02: 4.1459; 05: 3.9737; 06: 3.81; 07: 4.08; 08: 3.69 |  | 0
| 1 | 3 | 01: 4.05; 02: 4.0455; 05: 4.1727; 06: 3.81; 07: 4.2; 08: 3.67 |  | 16
| 1 | 4 | 01: 4.02; 02: 4.0644; 05: 4.1025; 06: 3.75; 07: 4.2; 08: 3.61 |  | 4.0351
| 1 | 5 | 01: 3.98; 02: 4.1298; 05: 4.1676; 06: 3.81; 07: 4.18; 08: 3.62 |  | 102
| 1 | 6 | 01: 4.09; 02: 4.1368; 05: 4.1908; 06: 3.69; 07: 4.17; 08: 3.57
| Mean |  | 01: 3.995; 02: 4.0992; 05: 4.0927; 06: 3.7833; 07: 4.1867; 08: 3.63
| Median |  | 01: 4; 02: 4.1012; 05: 4.1351; 06: 3.81; 07: 4.19; 08: 3.62
| Std Dev. |  | 01: 0.071764; 02: 0.043217; 05: 0.10634; 06: 0.053166; 07: 0.067429; 08: 0.043359
| Rel.Std.Dev. |  | 01: 1.796%; 02: 1.054%; 05: 2.598%; 06: 1.405%; 07: 1.611%; 08: 1.194%
| PDM3 |  | 01: -0.994%; 02: 1.587%; 05: 1.427%; 06: -6.24%; 07: 3.756%; 08: -10.04%
| Z-Score (Absolute) |  | 01: 0.33; 02: 0.38; 05: 0.33; 06: 1.65; 07: 0.97; 08: 2.82
| Analytical results for Pt in OREAS 600 (Indicative Value < 5 ppb) | [23 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab; 05: Lab |  | 1
| No | No | 01: 09; 02: 14; 05: 25 |  | ppb
|  |  | 01: AR*OES/MS; 02: AR*MS; 05: AR*MS |  | 0
|  |  | 01: 20g; 02: 0.5g; 05: 40g |  | 0
| 1 | 1 | 01: < 5; 02: < 2; 05: < 5 |  | 1
| 1 | 2 | 01: < 5; 02: < 2; 05: 5 |  | 6
| 1 | 3 | 01: < 5; 02: < 2; 05: < 5 |  | 16
| 1 | 4 | 01: < 5; 02: < 2; 05: < 5 |  | < 5
| 1 | 5 | 01: < 5; 02: < 2; 05: 5 |  | 12
| 1 | 6 | 01: < 5; 02: < 2; 05: < 5
| Mean |  | 05: 5
| Median |  | 05: 5
| Std Dev. |  | 05: 0
Rel.Std.Dev. | bar: 05
PDM3
Z-Score (Absolute) | bar: 01 to 05
| Analytical results for Rb in OREAS 600 (Indicative Value 10.8 ppm) | [23 empty] | Indicative
Rb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 10: 24; 11: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS |  | 1
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 10: 40g; 11: 40g |  | 1
| 1 | 1 | 01: 14.33; 02: 14.8569; 05: 10.7115; 06: 8.33; 07: 11.64; 08: 11.8; 09: 8.44; 10: 7.4; 11: 9.55 |  | 1
| 1 | 2 | 01: 14.65; 02: 14.286; 05: 10.5679; 06: 8.27; 07: 11.06; 08: 12.5; 09: 8.44; 10: 7.8; 11: 9.65 |  | 7
| 1 | 3 | 01: 14.63; 02: 14.4181; 05: 11.4276; 06: 8.21; 07: 10.96; 08: 10.4; 09: 8.37; 10: 7.7; 11: 9.45 |  | 16
| 1 | 4 | 01: 14.15; 02: 14.2275; 05: 10.844; 06: 8.1; 07: 11.32; 08: 11.7; 09: 8.27; 10: 8; 11: 9.05 |  | 10.7639
| 1 | 5 | 01: 14.38; 02: 14.8622; 05: 11.4249; 06: 8; 07: 11.76; 08: 12; 09: 8.13; 10: 7.4; 11: 9.1 |  | 13
| 1 | 6 | 01: 14.95; 02: 14.4732; 05: 11.2228; 06: 8.06; 07: 11.61; 08: 13; 09: 8.17; 10: 8.2; 11: 9
| Mean |  | 01: 14.515; 02: 14.5207; 05: 11.0331; 06: 8.1617; 07: 11.3917; 08: 11.9; 09: 8.3033; 10: 7.75; 11: 9.3
| Median |  | 01: 14.505; 02: 14.4457; 05: 11.0334; 06: 8.155; 07: 11.465; 08: 11.9; 09: 8.32; 10: 7.75; 11: 9.275
| Std Dev. |  | 01: 0.28508; 02: 0.27692; 05: 0.37434; 06: 0.12859; 07: 0.3306; 08: 0.88091; 09: 0.13471; 10: 0.32094; 11: 0.28284
| Rel.Std.Dev. |  | 01: 1.964%; 02: 1.907%; 05: 3.393%; 06: 1.576%; 07: 2.902%; 08: 7.403%; 09: 1.622%; 10: 4.141%; 11: 3.041%
| PDM3 |  | 01: 34.85%; 02: 34.9%; 05: 2.501%; 06: -24.18%; 07: 5.832%; 08: 10.55%; 09: -22.86%; 10: -28%; 11: -13.6%
| Z-Score (Absolute) |  | 01: 0.86; 02: 0.86; 05: 0; 06: 0.71; 07: 0.09; 08: 0.21; 09: 0.67; 10: 0.81; 11: 0.43
| Analytical results for Re in OREAS 600 (Indicative Value < 1 ppb) | [23 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 14; 07: 15 |  | ppb
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*OES; 06: AR*MS; 07: AR*OES/MS |  | 0
|  |  | 01: 40g; 02: 20g; 05: 0.15g; 06: 0.5g; 07: 01g |  | 0
| 1 | 1 | 01: < 1; 02: < 1; 05: < 5000; 06: < 1; 07: < 10 |  | 1
| 1 | 2 | 01: 1; 02: < 1; 05: < 5000; 06: < 1; 07: < 10 |  | 8
| 1 | 3 | 01: < 1; 02: < 1; 05: < 5000; 06: < 1; 07: < 10 |  | 16
| 1 | 4 | 01: < 1; 02: < 1; 05: < 5000; 06: < 1; 07: < 10 |  | < 1
| 1 | 5 | 01: < 1; 02: < 1; 05: < 5000; 06: < 1; 07: < 10 |  | 14
| 1 | 6 | 01: < 1; 02: < 1; 05: < 5000; 06: < 1; 07: < 10
Mean | bar: 01
Median | bar: 01
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 07
| Analytical results for Ru in OREAS 600 (Indicative Value < 2 ppb) | [23 empty] | Indicative
Ru | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 14 |  | ppb
|  |  | 01: AR*MS |  | 0
|  |  | 01: 0.5g |  | 0
| 1 | 1 | 01: < 2 |  | 1
| 1 | 2 | 01: < 2 |  | 9
| 1 | 3 | 01: < 2 |  | 16
| 1 | 4 | 01: < 2 |  | < 2
| 1 | 5 | 01: < 2 |  | 15
| 1 | 6 | 01: < 2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for S in OREAS 600 (Certified Value 1.66 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 19 | 23 | 24 | 25 |  |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  |  |  | 3
| 1 | 1 | 1.7833 | 1.5451 | 1.67 | 1.73 | 1.7 | 1.76 | 1.659 | 1.86 | 1.63 | 1.6699 | 1.6576 | 1.581 | 1.55 | - | 1.628 | 1.61 | 1.5 | 1.61 |  |  |  |  | 1
| 1 | 2 | 1.7767 | 1.6539 | 1.74 | 1.72 | 1.67 | 1.72 | 1.6399 | 1.94 | 1.6 | 1.6976 | 1.6082 | 1.596 | 1.57 | - | 1.668 | 1.59 | 1.67 | 1.64 |  |  |  |  | 0
| 1 | 3 | 1.7933 | 1.6312 | 1.72 | 1.73 | 1.68 | 1.72 | 1.6548 | 1.95 | 1.61 | 1.7023 | 1.6571 | 1.592 | 1.55 | - | 1.624 | 1.59 | 1.67 | 1.64 |  |  |  |  | 16
| 1 | 4 | 1.7767 | 1.5779 | 1.75 | 1.74 | 1.7 | 1.7 | 1.6887 | 1.7 | 1.62 | 1.6861 | 1.6256 | 1.595 | 1.56 | - | 1.647 | 1.61 | 1.66 | 1.68 |  |  |  |  | 1.6587
| 1 | 5 | 1.79 | 1.5971 | 1.7 | 1.73 | 1.67 | 1.71 | 1.6729 | 1.92 | 1.62 | 1.6555 | 1.6625 | 1.606 | 1.57 | - | 1.722 | 1.59 | 1.59 | 1.61 |  |  |  |  | 103
| 1 | 6 | 1.7667 | 1.6004 | 1.75 | 1.71 | 1.71 | 1.7 | 1.6346 | 1.9 | 1.62 | 1.6997 | 1.6247 | 1.603 | 1.57 | - | 1.662 | 1.58 | 1.71 | 1.61
| Mean |  | 1.7811 | 1.6009 | 1.7217 | 1.7267 | 1.6883 | 1.7183 | 1.6583 | 1.8783 | 1.6167 | 1.6852 | 1.6393 | 1.5955 | 1.5617 |  | 1.6585 | 1.595 | 1.6333 | 1.6317
| Median |  | 1.78 | 1.5987 | 1.73 | 1.73 | 1.69 | 1.715 | 1.6569 | 1.91 | 1.62 | 1.6918 | 1.6414 | 1.5955 | 1.565 |  | 1.6545 | 1.59 | 1.665 | 1.625
| Std Dev. |  | 0.0098131 | 0.03843 | 0.031885 | 0.010328 | 0.017224 | 0.022286 | 0.020271 | 0.093041 | 0.010328 | 0.018828 | 0.022619 | 0.0088261 | 0.0098319 |  | 0.035742 | 0.012247 | 0.07607 | 0.027869
| Rel.Std.Dev. |  | 0.551% | 2.4% | 1.852% | 0.5981% | 1.02% | 1.297% | 1.222% | 4.953% | 0.6388% | 1.117% | 1.38% | 0.5532% | 0.6296% |  | 2.155% | 0.7679% | 4.657% | 1.708%
| PDM3 |  | 7.381% | -3.481% | 3.798% | 4.099% | 1.788% | 3.597% | -0.02224% | 13.24% | -2.533% | 1.598% | -1.169% | -3.809% | -5.849% |  | -0.01065% | -3.839% | -1.528% | -1.628%
| Z-Score (Absolute) |  | 1.44 | 0.67 | 0.74 | 0.8 | 0.35 | 0.7 | 0 | 2.99 | 0.49 | 0.31 | 0.23 | 0.74 | 1.13 | NA | 0 | 0.74 | 0.02 | 0.31
| Analytical results for Sb in OREAS 600 (Certified Value < 15 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  | 1
| 1 | 1 | 5.1543 | 10.18 | 10 | 7 | 8 | 12 | 3.8297 | 10 | 5 | 12 | - | 10.93 | 12.13 | 13 | 1.53 | < 5 | 7.95 | 7 | < 3 | 5.9 |  |  | 1
| 1 | 2 | 5.103 | 9.9 | 12 | 7 | 8 | 14 | 4.0744 | < 10 | < 5 | 9 | - | 10.89 | 11.57 | 16 | 1.42 | 7 | 8.345 | 5 | 4 | 6.1 |  |  | 0
| 1 | 3 | 5.201 | 9.83 | 12 | 7 | 6 | 16 | 3.2076 | < 10 | 5 | 6 | - | 10.53 | 10.99 | 15 | 1.21 | 6 | 8.572 | 5 | 4 | 5.9 |  |  | 16
| 1 | 4 | 4.7763 | 9.55 | 10 | 7 | 7 | 13 | 3.346 | 10 | < 5 | 6 | - | 10.11 | 12.01 | 12 | 1.34 | < 5 | 8.238 | 6 | 4 | 6.6 |  |  | < 15
| 1 | 5 | 5.32 | 9.46 | 10 | 7 | 9 | 15 | 4.285 | 10 | < 5 | 9 | - | 9.73 | 12.09 | 14 | 1.29 | 8 | 8.693 | 8 | < 3 | 6.8 |  |  | 104
| 1 | 6 | 5.306 | 10.32 | 11 | 6 | 7 | 14 | 5.1582 | < 10 | < 5 | 7 | - | 10.56 | 11.81 | 14 | 1.47 | < 5 | 7.624 | 7 | 3 | 7.9
| Mean |  | 5.1434 | 9.8733 | 10.8333 | 6.8333 | 7.5 | 14 | 3.9835 | 10 | 5 | 8.1667 |  | 10.4583 | 11.7667 | 14 | 1.3767 | 7 | 8.237 | 6.3333 | 3.75 | 6.5333
| Median |  | 5.1777 | 9.865 | 10.5 | 7 | 7.5 | 14 | 3.952 | 10 | 5 | 8 |  | 10.545 | 11.91 | 14 | 1.38 | 7 | 8.2915 | 6.5 | 4 | 6.35
| Std Dev. |  | 0.19871 | 0.33809 | 0.98319 | 0.40825 | 1.0488 | 1.4142 | 0.70867 | 0 | 0 | 2.3166 |  | 0.46391 | 0.43348 | 1.4142 | 0.11894 | 1 | 0.39748 | 1.2111 | 0.5 | 0.76594
| Rel.Std.Dev. |  | 3.863% | 3.424% | 9.076% | 5.974% | 13.98% | 10.1% | 17.79% | 0% | 0% | 28.37% |  | 4.436% | 3.684% | 10.1% | 8.64% | 14.29% | 4.826% | 19.12% | 13.33% | 11.72%
PDM3
| Z-Score (Absolute) |  | 0.62 | 0.64 | 0.9 | 0.18 | 0 | 1.75 | 0.95 | 0.67 | 0.67 | 0.18 | NA | 0.8 | 1.15 | 1.75 | 1.65 | 0.13 | 0.2 | 0.31 | 1.01 | 0.26
| Analytical results for Sc in OREAS 600 (Certified Value 2.06 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g | 40g |  |  |  | 2
| 1 | 1 | 2.2367 | 2.4 | 2 | 2 | 2 | 2 | 2.4851 | 2 | 2 | 1.7957 | 2 | 2 | 1.5 | 2.7 | < 3 | 2.11 | < 5 | 2 | 2 |  |  |  | 1
| 1 | 2 | 2.23 | 2.3 | 2 | 2 | 2 | 2 | 2.4504 | 2 | 2 | 1.7683 | 2 | 2 | 1.5 | 2.6 | < 3 | 2.19 | < 5 | 2 | 2 |  |  |  | 0
| 1 | 3 | 2.2133 | 2.3 | 2 | 2 | 2 | 2 | 2.4002 | 2 | 2 | 1.8173 | 2 | 2 | 1.4 | 2.1 | < 3 | 2.27 | < 5 | 2 | 2 |  |  |  | 16
| 1 | 4 | 2.22 | 2.2 | 2 | 2 | 2 | 2 | 2.5618 | 2 | 2 | 1.8093 | 2 | 2 | 1.5 | 2.8 | < 3 | 2.2 | < 5 | 3 | 2 |  |  |  | 2.0607
| 1 | 5 | 2.2133 | 2.2 | 2 | 2 | 2 | 2 | 2.6056 | 2 | 2 | 1.8606 | 2 | 2 | 1.5 | 2.4 | < 3 | 2.26 | < 5 | 2 | 2 |  |  |  | 105
| 1 | 6 | 2.2 | 2.3 | 2 | 2 | 2 | 2 | 2.4796 | 2 | 2 | 1.8566 | 2 | 2 | 1.5 | 2.5 | < 3 | 2.16 | < 5 | 3 | 2
| Mean |  | 2.2189 | 2.2833 | 2 | 2 | 2 | 2 | 2.4971 | 2 | 2 | 1.818 | 2 | 2 | 1.4833 | 2.5167 |  | 2.1983 |  | 2.3333 | 2
| Median |  | 2.2167 | 2.3 | 2 | 2 | 2 | 2 | 2.4823 | 2 | 2 | 1.8133 | 2 | 2 | 1.5 | 2.55 |  | 2.195 |  | 2 | 2
| Std Dev. |  | 0.013109 | 0.075277 | 0 | 0 | 0 | 0 | 0.074822 | 0 | 0 | 0.035619 | 0 | 0 | 0.040825 | 0.24833 |  | 0.06047 |  | 0.5164 | 0
| Rel.Std.Dev. |  | 0.5908% | 3.297% | 0% | 0% | 0% | 0% | 2.996% | 0% | 0% | 1.959% | 0% | 0% | 2.752% | 9.867% |  | 2.751% |  | 22.13% | 0%
| PDM3 |  | 7.675% | 10.8% | -2.947% | -2.947% | -2.947% | -2.947% | 21.18% | -2.947% | -2.947% | -11.78% | -2.947% | -2.947% | -28.02% | 22.13% |  | 6.678% |  | 13.23% | -2.947%
Z-Score (Absolute) | bar: 01 to 25
| Analytical results for Se in OREAS 600 (Certified Value 6.77 ppm) | [23 empty] | Certified
Se | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 14; 08: 15; 09: 17; 11: 19; 12: 21; 13: 24; 14: 25 |  |  |  | ppm
|  |  | 01: AR*OES; 02: AR*MS; 05: AR*OES/MS; 06: AR*OES; 07: AR*MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*OES/MS; 12: AR*OES; 13: AR*MS; 14: AR*MS |  |  |  | 2
|  |  | 01: 0.5g; 02: 40g; 05: 20g; 06: 0.15g; 07: 0.5g; 08: 01g; 09: 0.25g; 11: 0.25g; 12: 0.25g; 13: 40g; 14: 40g |  |  |  | 2
| 1 | 1 | 01: 6.9167; 02: 7; 05: 6.9446; 06: 7; 07: 6.1; 08: 8; 09: 7; 11: 5.573; 12: < 5; 13: 6.5; 14: 6 |  |  |  | 1
| 1 | 2 | 01: 6.43; 02: 8; 05: 6.9798; 06: 7; 07: 6.1; 08: 7.78; 09: 7; 11: 5.562; 12: 7; 13: 7.5; 14: 6 |  |  |  | 0
| 1 | 3 | 01: 5.93; 02: 8; 05: 6.9803; 06: 7; 07: 6.2; 08: 7.66; 09: 6; 11: 5.686; 12: < 5; 13: 7; 14: 6 |  |  |  | 16
| 1 | 4 | 01: 6.7133; 02: 8; 05: 7.0042; 06: 8; 07: 5.9; 08: 7.91; 09: 7; 11: 5.534; 12: < 5; 13: 7.5; 14: 6 |  |  |  | 6.7749
| 1 | 5 | 01: 6.6; 02: 8; 05: 6.8696; 06: 5; 07: 6.2; 08: 7.64; 09: 7; 11: 5.372; 12: < 5; 13: 7.5; 14: 6 |  |  |  | 106
| 1 | 6 | 01: 6.5433; 02: 8; 05: 7.3373; 06: 6; 07: 6.1; 08: 8; 09: 7; 11: 5.078; 12: < 5; 13: 7.5; 14: 6
| Mean |  | 01: 6.5222; 02: 7.8333; 05: 7.0193; 06: 6.6667; 07: 6.1; 08: 7.8317; 09: 6.8333; 11: 5.4675; 12: 7; 13: 7.25; 14: 6
| Median |  | 01: 6.5717; 02: 8; 05: 6.98; 06: 7; 07: 6.1; 08: 7.845; 09: 7; 11: 5.548; 12: 7; 13: 7.5; 14: 6
| Std Dev. |  | 01: 0.33402; 02: 0.40825; 05: 0.16273; 06: 1.0328; 07: 0.10954; 08: 0.16229; 09: 0.40825; 11: 0.2159; 13: 0.41833; 14: 0
| Rel.Std.Dev. |  | 01: 5.121%; 02: 5.212%; 05: 2.318%; 06: 15.49%; 07: 1.796%; 08: 2.072%; 09: 5.974%; 11: 3.949%; 13: 5.77%; 14: 0%
| PDM3 |  | 01: -3.73%; 02: 15.62%; 05: 3.607%; 06: -1.598%; 07: -9.962%; 08: 15.6%; 09: 0.8623%; 11: -19.3%; 12: 3.322%; 13: 7.012%; 14: -11.44%
| Z-Score (Absolute) |  | 01: 0.5; 02: 1.62; 05: 0.3; 06: 0.27; 07: 1.19; 08: 1.62; 09: 0; 11: 2.21; 12: 0.27; 13: 0.67; 14: 1.35
| Analytical results for Si in OREAS 600 (Indicative Value 0.037 wt.%) | [23 empty] | Indicative
Si | Round | Replicate | 01: Lab |  |  |  | 1
| No | No | 01: 21 |  |  |  | wt.%
|  |  | 01: AR*OES |  |  |  | 3
|  |  | 01: 0.25g |  |  |  | 3
| 1 | 1 | 01: 0.03 |  |  |  | 1
| 1 | 2 | 01: 0.03 |  |  |  | 6
| 1 | 3 | 01: 0.04 |  |  |  | 16
| 1 | 4 | 01: 0.04 |  |  |  | 0.036667
| 1 | 5 | 01: 0.04 |  |  |  | 12
| 1 | 6 | 01: 0.04
| Mean |  | 01: 0.036667
| Median |  | 01: 0.04
| Std Dev. |  | 01: 0.005164
| Rel.Std.Dev. |  | 01: 14.08%
| PDM3 |  | 01: -7.772e-14%
Z-Score (Absolute) | bar: 01
| Analytical results for Sm in OREAS 600 (Certified Value 2.49 ppm) | [23 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  |  |  | 3
| 1 | 1 | 01: 2.46; 02: 2.79; 05: 2.1835; 06: 2.33; 07: 2.74; 08: 2.42 |  |  |  | 1
| 1 | 2 | 01: 2.54; 02: 2.7902; 05: 2.2746; 06: 2.36; 07: 2.58; 08: 2.4 |  |  |  | 0
| 1 | 3 | 01: 2.62; 02: 2.7441; 05: 2.2798; 06: 2.35; 07: 2.65; 08: 2.49 |  |  |  | 16
| 1 | 4 | 01: 2.56; 02: 2.6431; 05: 2.2027; 06: 2.38; 07: 2.65; 08: 2.39 |  |  |  | 2.4886
| 1 | 5 | 01: 2.54; 02: 2.7139; 05: 2.2299; 06: 2.33; 07: 2.62; 08: 2.44 |  |  |  | 107
| 1 | 6 | 01: 2.58; 02: 2.7789; 05: 2.2289; 06: 2.32; 07: 2.65; 08: 2.33
| Mean |  | 01: 2.55; 02: 2.7434; 05: 2.2332; 06: 2.345; 07: 2.6483; 08: 2.4117
| Median |  | 01: 2.55; 02: 2.7615; 05: 2.2294; 06: 2.34; 07: 2.65; 08: 2.41
| Std Dev. |  | 01: 0.053292; 02: 0.057544; 05: 0.038254; 06: 0.022583; 07: 0.052694; 08: 0.053448
| Rel.Std.Dev. |  | 01: 2.09%; 02: 2.098%; 05: 1.713%; 06: 0.963%; 07: 1.99%; 08: 2.216%
| PDM3 |  | 01: 2.467%; 02: 10.24%; 05: -10.26%; 06: -5.77%; 07: 6.419%; 08: -3.091%
| Z-Score (Absolute) |  | 01: 0.31; 02: 1.17; 05: 1.1; 06: 0.6; 07: 0.74; 08: 0.31
| Analytical results for Sn in OREAS 600 (Certified Value 1.18 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 14 | 15 | 16 | 17 | 18 | 19 | 24 | 25 | [10 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 40g | 40g | [10 empty] | 2
| 1 | 1 | < 5 | 1.34 | 1.2457 | < 10 | 1.2 | 1.32 | < 10 | 1.8 | < 10 | 1.044 | 1 | 1.2 | [10 empty] | 1
| 1 | 2 | < 5 | 1.35 | 1.2365 | < 10 | 1.2 | 1.21 | < 10 | 1.7 | < 10 | 1.034 | 1 | 1.2 | [10 empty] | 0
| 1 | 3 | < 5 | 1.31 | 1.263 | < 10 | 1.3 | 1.32 | < 10 | 1.4 | < 10 | 1.003 | 1 | 1.2 | [10 empty] | 16
| 1 | 4 | < 5 | 1.28 | 1.2831 | < 10 | 1.2 | 1.31 | < 10 | 1.5 | < 10 | 1 | 1 | 1.2 | [10 empty] | 1.179
| 1 | 5 | < 5 | 1.28 | 1.2668 | < 10 | 1.2 | 1.23 | < 10 | 1.5 | < 10 | 1.008 | 1 | 1.2 | [10 empty] | 108
| 1 | 6 | < 5 | 1.32 | 1.2554 | < 10 | 1.2 | 1.25 | < 10 | 1.6 | < 10 | 0.96 | 1 | 1.1
| Mean |  | 02: 1.3133; 05: 1.2584; 07: 1.2167; 08: 1.2733; 11: 1.5833; 13: 1.0082; 14: 1; 15: 1.1833
| Median |  | 02: 1.315; 05: 1.2592; 07: 1.2; 08: 1.28; 11: 1.55; 13: 1.0055; 14: 1; 15: 1.2
| Std Dev. |  | 02: 0.029439; 05: 0.016465; 07: 0.040825; 08: 0.049261; 11: 0.1472; 13: 0.029546; 14: 0; 15: 0.040825
| Rel.Std.Dev. |  | 02: 2.242%; 05: 1.308%; 07: 3.355%; 08: 3.869%; 11: 9.297%; 13: 2.931%; 14: 0%; 15: 3.45%
| PDM3 |  | 02: 11.39%; 05: 6.732%; 07: 3.192%; 08: 7.998%; 11: 34.29%; 13: -14.49%; 14: -15.18%; 15: 0.3647%
| Z-Score (Absolute) |  | NA | 0.79 | 0.22 | NA | 0.22 | 0.37 | NA | 3.59 | NA | 2.38 | 2.46 | 0.56
| Analytical results for Sr in OREAS 600 (Certified Value 38.4 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  | 1
| 1 | 1 | 34.4517 | 47.25 | 38 | 40 | 40 | 40 | 44.7013 | 43 | 44 | 38.2137 | 36.36 | 40 | 34.5 | 40.1 | 32 | 32.75 | 37 | 34 | 31 | 36.4 |  |  | 1
| 1 | 2 | 34.3397 | 47.91 | 40 | 39 | 40 | 39 | 43.4748 | 42 | 42 | 36.9276 | 36.53 | 41 | 35.2 | 39.9 | 33 | 33.28 | 38 | 34 | 35 | 37.9 |  |  | 0
| 1 | 3 | 34.566 | 47.9 | 41 | 39 | 39 | 39 | 43.354 | 43 | 42 | 38.1526 | 36.61 | 41 | 34.4 | 36.7 | 32 | 34.19 | 38 | 33 | 33 | 37.6 |  |  | 16
| 1 | 4 | 33.8403 | 47.54 | 40 | 40 | 40 | 38 | 43.8246 | 43 | 43 | 37.8111 | 35.79 | 40 | 35.4 | 40.7 | 33 | 33.63 | 39 | 34 | 32 | 37 |  |  | 38.3771
| 1 | 5 | 34.6243 | 48.21 | 40 | 39 | 39 | 39 | 44.7156 | 41 | 43 | 38.6734 | 36.46 | 41 | 35.2 | 39.7 | 33 | 34.38 | 40 | 34 | 33 | 36.9 |  |  | 109
| 1 | 6 | 34.3863 | 48.49 | 41 | 39 | 40 | 39 | 43.9191 | 43 | 43 | 37.4698 | 36.08 | 42 | 35.5 | 38.6 | 32 | 33.25 | 40 | 33 | 32 | 36.4
| Mean |  | 34.3681 | 47.8833 | 40 | 39.3333 | 39.6667 | 39 | 43.9982 | 42.5 | 42.8333 | 37.8747 | 36.305 | 40.8333 | 35.0333 | 39.2833 | 32.5 | 33.58 | 38.6667 | 33.6667 | 32.6667 | 37.0333
| Median |  | 34.419 | 47.905 | 40 | 39 | 40 | 39 | 43.8719 | 43 | 43 | 37.9819 | 36.41 | 41 | 35.2 | 39.8 | 32.5 | 33.455 | 38.5 | 34 | 32.5 | 36.95
| Std Dev. |  | 0.27981 | 0.44603 | 1.0954 | 0.5164 | 0.5164 | 0.63246 | 0.58893 | 0.83666 | 0.75277 | 0.61564 | 0.31195 | 0.75277 | 0.46762 | 1.44 | 0.54772 | 0.61683 | 1.2111 | 0.5164 | 1.3663 | 0.61536
| Rel.Std.Dev. |  | 0.8142% | 0.9315% | 2.739% | 1.313% | 1.302% | 1.622% | 1.339% | 1.969% | 1.757% | 1.625% | 0.8592% | 1.844% | 1.335% | 3.666% | 1.685% | 1.837% | 3.132% | 1.534% | 4.182% | 1.662%
| PDM3 |  | -10.45% | 24.77% | 4.229% | 2.492% | 3.36% | 1.623% | 14.65% | 10.74% | 11.61% | -1.309% | -5.399% | 6.4% | -8.713% | 2.361% | -15.31% | -12.5% | 0.7544% | -12.27% | -14.88% | -3.502%
| Z-Score (Absolute) |  | 0.97 | 1.97 | 0.25 | 0.11 | 0.18 | 0.04 | 1.12 | 0.8 | 0.87 | 0.21 | 0.55 | 0.44 | 0.83 | 0.21 | 1.38 | 1.14 | 0.04 | 1.12 | 1.34 | 0.39
| Analytical results for Ta in OREAS 600 (Indicative Value 0.12 ppm) | [23 empty] | Indicative
Ta | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 14; 06: 15; 07: 17; 08: 19; 09: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*MS |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 0.5g; 06: 01g; 07: 0.25g; 08: 0.25g; 09: 40g |  |  | 2
| 1 | 1 | 01: 0.14; 02: 0.0074329; 05: < 0.01; 06: < 0.1; 07: 0.33; 08: < 0.05; 09: < 0.005 |  |  | 1
| 1 | 2 | 01: 0.14; 02: 0.0068659; 05: < 0.01; 06: < 0.1; 07: 0.22; 08: < 0.05; 09: < 0.005 |  |  | 7
| 1 | 3 | 01: 0.15; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.2; 08: < 0.05; 09: < 0.005 |  |  | 16
| 1 | 4 | 01: 0.13; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.21; 08: < 0.05; 09: < 0.005 |  |  | 0.12024
| 1 | 5 | 01: 0.12; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.18; 08: < 0.05; 09: < 0.005 |  |  | 13
| 1 | 6 | 01: 0.13; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.18; 08: < 0.05; 09: < 0.005
| Mean |  | 01: 0.135; 02: 0.0071494; 07: 0.22
| Median |  | 01: 0.135; 02: 0.0071494; 07: 0.205
| Std Dev. |  | 01: 0.010488; 02: 0.00040097; 07: 0.056214
| Rel.Std.Dev. |  | 01: 7.769%; 02: 5.608%; 07: 25.55%
| PDM3 |  | 01: 12.28%; 02: -94.05%; 07: 82.97%
| Z-Score (Absolute) |  | 01: 0; 02: 1.03; 05: NA; 06: NA; 07: 0.67; 08: NA; 09: NA
| Analytical results for Tb in OREAS 600 (Certified Value 0.25 ppm) | [23 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 11: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*MS |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 11: 40g |  |  | 3
| 1 | 1 | 01: 0.25; 02: 0.28704; 05: 0.23509; 06: 0.258; 07: 0.29; 08: 0.23; 09: 0.244; 11: 0.235 |  |  | 1
| 1 | 2 | 01: 0.25; 02: 0.29369; 05: 0.2337; 06: 0.258; 07: 0.27; 08: 0.24; 09: 0.254; 11: 0.24 |  |  | 0
| 1 | 3 | 01: 0.26; 02: 0.27858; 05: 0.24087; 06: 0.265; 07: 0.28; 08: 0.19; 09: 0.255; 11: 0.235 |  |  | 16
| 1 | 4 | 01: 0.26; 02: 0.28445; 05: 0.22816; 06: 0.253; 07: 0.28; 08: 0.24; 09: 0.253; 11: 0.22 |  |  | 0.25423
| 1 | 5 | 01: 0.26; 02: 0.29245; 05: 0.2735; 06: 0.257; 07: 0.28; 08: 0.22; 09: 0.253; 11: 0.22 |  |  | 110
| 1 | 6 | 01: 0.26; 02: 0.28942; 05: 0.25065; 06: 0.258; 07: 0.29; 08: 0.23; 09: 0.241; 11: 0.23
| Mean |  | 01: 0.25667; 02: 0.2876; 05: 0.24366; 06: 0.25817; 07: 0.28167; 08: 0.225; 09: 0.25; 11: 0.23
| Median |  | 01: 0.26; 02: 0.28823; 05: 0.23798; 06: 0.258; 07: 0.28; 08: 0.23; 09: 0.253; 11: 0.2325
| Std Dev. |  | 01: 0.005164; 02: 0.0055772; 05: 0.016491; 06: 0.0038687; 07: 0.0075277; 08: 0.018708; 09: 0.005933; 11: 0.0083666
| Rel.Std.Dev. |  | 01: 2.012%; 02: 1.939%; 05: 6.768%; 06: 1.499%; 07: 2.673%; 08: 8.315%; 09: 2.373%; 11: 3.638%
| PDM3 |  | 01: 0.9604%; 02: 13.13%; 05: -4.154%; 06: 1.55%; 07: 10.79%; 08: -11.5%; 09: -1.662%; 11: -9.529%
| Z-Score (Absolute) |  | 01: 0.12; 02: 1.25; 05: 0.57; 06: 0.18; 07: 1.03; 08: 0.78; 09: 0.12; 11: 0.85
| Analytical results for Te in OREAS 600 (Certified Value 6.8 ppm) | [23 empty] | Certified
Te | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 14; 07: 15; 08: 17; 09: 19; 11: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*OES; 06: AR*MS; 07: AR*OES; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*MS |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 0.15g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 11: 40g |  |  | 2
| 1 | 1 | 01: 5.31; 02: 6.2112; 05: 6; 06: 6.89; 07: 9; 08: 3.44; 09: 6.99; 11: 6.6 |  |  | 1
| 1 | 2 | 01: 6.08; 02: 5.9232; 05: 7; 06: 6.9; 07: 7; 08: 4.61; 09: 7.282; 11: 6.74 |  |  | 0
| 1 | 3 | 01: 6.84; 02: 6.0429; 05: 7; 06: 6.94; 07: 8; 08: 4.27; 09: 7.19; 11: 6.86 |  |  | 16
| 1 | 4 | 01: 6.96; 02: 6.1463; 05: 6; 06: 6.9; 07: 8; 08: 4.98; 09: 7.318; 11: 6.74 |  |  | 6.7983
| 1 | 5 | 01: 7.04; 02: 6.0338; 05: 5; 06: 6.88; 07: 7; 08: 5.14; 09: 7.234; 11: 6.86 |  |  | 111
| 1 | 6 | 01: 7.75; 02: 6.0942; 05: 6; 06: 6.89; 07: 8; 08: 5.41; 09: 7.061; 11: 6.82
| Mean |  | 01: 6.6633; 02: 6.0753; 05: 6.1667; 06: 6.9; 07: 7.8333; 08: 4.6417; 09: 7.1792; 11: 6.77
| Median |  | 01: 6.9; 02: 6.0685; 05: 6; 06: 6.895; 07: 8; 08: 4.795; 09: 7.212; 11: 6.78
| Std Dev. |  | 01: 0.85015; 02: 0.099774; 05: 0.75277; 06: 0.020976; 07: 0.75277; 08: 0.71222; 09: 0.12862; 11: 0.099398
| Rel.Std.Dev. |  | 01: 12.76%; 02: 1.642%; 05: 12.21%; 06: 0.304%; 07: 9.61%; 08: 15.34%; 09: 1.792%; 11: 1.468%
| PDM3 |  | 01: -1.985%; 02: -10.63%; 05: -9.29%; 06: 1.497%; 07: 15.23%; 08: -31.72%; 09: 5.603%; 11: -0.4156%
| Z-Score (Absolute) |  | 01: 0.07; 02: 0.85; 05: 0.73; 06: 0.24; 07: 1.49; 08: 2.76; 09: 0.62; 11: 0.07
| Analytical results for Th in OREAS 600 (Certified Value 4.91 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 08 | 09 | 13 | 14 | 15 | 17 | 18 | 19 | 24 | 25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [8 empty] | 2
|  |  | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 40g | 40g | [8 empty] | 3
| 1 | 1 | 4.71 | < 20 | < 20 | < 20 | < 20 | 4.9279 | 5.7048 | 4.98 | 5.51 | 4.7 | < 10 | 4.939 | 4.15 | 4.66 | [8 empty] | 1
| 1 | 2 | 4.8 | < 20 | < 20 | < 20 | < 20 | 4.7968 | 5.967 | 5.08 | 5.16 | 5.1 | < 10 | 4.942 | 4.5 | 4.8 | [8 empty] | 0
| 1 | 3 | 4.96 | < 20 | < 20 | < 20 | < 20 | 4.8229 | 5.8594 | 5.09 | 5.23 | 4.3 | < 10 | 5.123 | 4.3 | 4.8 | [8 empty] | 16
| 1 | 4 | 4.96 | < 20 | < 20 | < 20 | < 20 | 4.9662 | 5.8674 | 4.96 | 5.46 | 5 | < 10 | 5.192 | 4.4 | 4.7 | [8 empty] | 4.9082
| 1 | 5 | 5.02 | < 20 | < 20 | < 20 | < 20 | 4.8716 | 5.9244 | 5.03 | 5.29 | 5 | < 10 | 5.281 | 4.4 | 4.72 | [8 empty] | 112
| 1 | 6 | 5.09 | < 20 | < 20 | < 20 | < 20 | 4.9806 | 5.8987 | 4.93 | 5.23 | 5.3 | < 10 | 5.303 | 4.45 | 4.68
| Mean |  | 01: 4.9233; 08: 4.8943; 09: 5.8703; 11: 5.0117; 12: 5.3133; 13: 4.9; 15: 5.13; 16: 4.3667; 17: 4.7267
| Median |  | 01: 4.96; 08: 4.8997; 09: 5.883; 11: 5.005; 12: 5.26; 13: 5; 15: 5.1575; 16: 4.4; 17: 4.71
| Std Dev. |  | 01: 0.1418; 08: 0.076008; 09: 0.090101; 11: 0.065549; 12: 0.1401; 13: 0.35214; 15: 0.16028; 16: 0.12517; 17: 0.060222
| Rel.Std.Dev. |  | 01: 2.88%; 08: 1.553%; 09: 1.535%; 11: 1.308%; 12: 2.637%; 13: 7.186%; 15: 3.124%; 16: 2.866%; 17: 1.274%
| PDM3 |  | 01: 0.3073%; 08: -0.2837%; 09: 19.6%; 11: 2.107%; 12: 8.253%; 13: -0.1681%; 15: 4.518%; 16: -11.03%; 17: -3.7%
| Z-Score (Absolute) |  | 0 | NA | NA | NA | NA | 0.1 | 3.25 | 0.3 | 1.34 | 0.08 | NA | 0.71 | 1.91 | 0.67
| Analytical results for Ti in OREAS 600 (Certified Value < 0.01 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  |  |  | 3
| 1 | 1 | < 0.01 | 0.0049 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.0044975 | < 0.01 | < 0.01 | 0.0028 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | < 0.005 |  |  |  |  | 1
| 1 | 2 | < 0.01 | 0.005 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.0043627 | < 0.01 | < 0.01 | 0.0027 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | < 0.005 |  |  |  |  | 0
| 1 | 3 | < 0.01 | 0.0052 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.0042045 | < 0.01 | < 0.01 | 0.0026 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | < 0.005 |  |  |  |  | 16
| 1 | 4 | < 0.01 | 0.005 | 0.01 | < 0.01 | < 0.01 | 0.01 | 0.0041898 | < 0.01 | < 0.01 | 0.0026 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | < 0.005 |  |  |  |  | < 0.01
| 1 | 5 | < 0.01 | 0.0048 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.0047913 | < 0.01 | < 0.01 | 0.0026 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | < 0.005 |  |  |  |  | 113
| 1 | 6 | < 0.01 | 0.0052 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.0044416 | < 0.01 | < 0.01 | 0.0025 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | < 0.005
| Mean |  | 02: 0.0050167; 05: 0.01; 08: 0.01; 09: 0.0044146; 13: 0.0026333
| Median |  | 02: 0.005; 05: 0.01; 08: 0.01; 09: 0.0044022; 13: 0.0026
| Std Dev. |  | 02: 0.00016021; 05: 0; 08: 0; 09: 0.0002222; 13: 0.00010328
| Rel.Std.Dev. |  | 02: 3.194%; 05: 0%; 08: 0%; 09: 5.033%; 13: 3.922%
PDM3
| Z-Score (Absolute) |  | NA | 0 | 0.67 | NA | NA | 2.04 | 0.25 | NA | NA | 0.99 | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tl in OREAS 600 (Certified Value 0.56 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  |  | 3
| 1 | 1 | < 5 | 0.61 | < 10 | < 10 | < 10 | < 10 | 0.48992 | < 5 | < 10 | 0.56993 | 0.54 | 0.62 | 0.59 | < 10 | 0.53 | < 2 | < 5 | 0.5 | 0.5 |  |  |  | 1
| 1 | 2 | < 5 | 0.62 | < 10 | < 10 | < 10 | < 10 | 0.45852 | < 5 | < 10 | 0.55807 | 0.54 | 0.63 | 0.64 | < 10 | 0.541 | < 2 | < 5 | 0.6 | 0.5 |  |  |  | 0
| 1 | 3 | < 5 | 0.64 | < 10 | < 10 | < 10 | < 10 | 0.4868 | < 5 | < 10 | 0.59949 | 0.54 | 0.62 | 0.55 | < 10 | 0.535 | < 2 | < 5 | 0.5 | 0.5 |  |  |  | 16
| 1 | 4 | < 5 | 0.61 | < 10 | < 10 | < 10 | < 10 | 0.49035 | < 5 | < 10 | 0.55406 | 0.53 | 0.61 | 0.58 | < 10 | 0.539 | < 2 | < 5 | 0.6 | 0.5 |  |  |  | 0.562
| 1 | 5 | < 5 | 0.62 | < 10 | < 10 | < 10 | < 10 | 0.48819 | < 5 | < 10 | 0.59532 | 0.54 | 0.63 | 0.63 | < 10 | 0.556 | < 2 | < 5 | 0.6 | 0.5 |  |  |  | 114
| 1 | 6 | < 5 | 0.64 | < 10 | < 10 | < 10 | < 10 | 0.48421 | < 5 | < 10 | 0.55069 | 0.53 | 0.62 | 0.66 | < 10 | 0.552 | < 2 | < 5 | 0.6 | 0.5
| Mean |  | 02: 0.62333; 09: 0.483; 13: 0.57126; 14: 0.53667; 15: 0.62167; 17: 0.60833; 19: 0.54217; 24: 0.56667; 25: 0.5
| Median |  | 02: 0.62; 09: 0.4875; 13: 0.564; 14: 0.54; 15: 0.62; 17: 0.61; 19: 0.54; 24: 0.6; 25: 0.5
| Std Dev. |  | 02: 0.013663; 09: 0.012198; 13: 0.02131; 14: 0.005164; 15: 0.0075277; 17: 0.041673; 19: 0.0099883; 24: 0.05164; 25: 0
| Rel.Std.Dev. |  | 02: 2.192%; 09: 2.526%; 13: 3.73%; 14: 0.9622%; 15: 1.211%; 17: 6.85%; 19: 1.842%; 24: 9.113%; 25: 0%
| PDM3 |  | 02: 10.91%; 09: -14.06%; 13: 1.648%; 14: -4.508%; 15: 10.62%; 17: 8.245%; 19: -3.529%; 24: 0.8306%; 25: -11.03%
| Z-Score (Absolute) |  | NA | 0.92 | NA | NA | NA | NA | 1.27 | NA | NA | 0.07 | 0.49 | 0.89 | 0.67 | NA | 0.4 | NA | NA | 0 | 1.08
| Analytical results for Tm in OREAS 600 (Indicative Value 0.064 ppm) | [23 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 14; 06: 15; 07: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS |  |  |  | 3
|  |  | 01: 40g; 02: 20g; 05: 0.5g; 06: 01g; 07: 40g |  |  |  | 3
| 1 | 1 | 01: 0.06; 02: 0.064077; 05: 0.07; 06: < 0.1; 07: 0.065 |  |  |  | 1
| 1 | 2 | 01: 0.06; 02: 0.068425; 05: 0.07; 06: < 0.1; 07: 0.065 |  |  |  | 4
| 1 | 3 | 01: 0.06; 02: 0.05669; 05: 0.07; 06: < 0.1; 07: 0.07 |  |  |  | 16
| 1 | 4 | 01: 0.06; 02: 0.058833; 05: 0.07; 06: < 0.1; 07: 0.07 |  |  |  | 0.064324
| 1 | 5 | 01: 0.06; 02: 0.060341; 05: 0.06; 06: < 0.1; 07: 0.065 |  |  |  | 14
| 1 | 6 | 01: 0.07; 02: 0.055418; 05: 0.07; 06: < 0.1; 07: 0.065
| Mean |  | 01: 0.061667; 02: 0.060631; 05: 0.068333; 07: 0.066667
| Median |  | 01: 0.06; 02: 0.059587; 05: 0.07; 07: 0.065
| Std Dev. |  | 01: 0.0040825; 02: 0.0048714; 05: 0.0040825; 07: 0.002582
| Rel.Std.Dev. |  | 01: 6.62%; 02: 8.034%; 05: 5.974%; 07: 3.873%
| PDM3 |  | 01: -4.132%; 02: -5.742%; 05: 6.232%; 07: 3.641%
| Z-Score (Absolute) |  | 01: 0.56; 02: 0.79; 05: 0.93; 06: NA; 07: 0.56
| Analytical results for U in OREAS 600 (Certified Value 1.02 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 05 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 24 | 25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [8 empty] | 2
|  |  | 40g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 40g | 40g | [8 empty] | 3
| 1 | 1 | 1.03 | < 10 | < 10 | < 10 | 1.0449 | < 10 | 0.93725 | 0.95 | 1.19 | 1.22 | < 10 | 1.003 | 0.9 | 0.92 | [8 empty] | 1
| 1 | 2 | 1.03 | < 10 | < 10 | < 10 | 0.99733 | < 10 | 0.95264 | 0.96 | 1.14 | 1.2 | < 10 | 0.992 | 1 | 0.92 | [8 empty] | 0
| 1 | 3 | 1.06 | < 10 | < 10 | < 10 | 1.0171 | < 10 | 0.96487 | 0.94 | 1.18 | 1.06 | < 10 | 1.049 | 0.95 | 0.92 | [8 empty] | 16
| 1 | 4 | 1.06 | < 10 | < 10 | < 10 | 1.0327 | < 10 | 0.94161 | 0.94 | 1.17 | 1.06 | < 10 | 1.058 | 0.9 | 0.9 | [8 empty] | 1.0248
| 1 | 5 | 1.07 | < 10 | < 10 | < 10 | 1.0456 | < 10 | 1.0093 | 0.95 | 1.17 | 1.17 | < 10 | 1.067 | 1.05 | 0.9 | [8 empty] | 115
| 1 | 6 | 1.11 | < 10 | < 10 | < 10 | 1.0478 | < 10 | 0.97871 | 0.93 | 1.15 | 1.26 | < 10 | 1.089 | 0.85 | 0.9
| Mean |  | 01: 1.06; 07: 1.0309; 09: 0.96406; 11: 0.945; 12: 1.1667; 13: 1.1617; 15: 1.043; 17: 0.94167; 18: 0.91
| Median |  | 01: 1.06; 07: 1.0388; 09: 0.95875; 11: 0.945; 12: 1.17; 13: 1.185; 15: 1.0535; 17: 0.925; 18: 0.91
| Std Dev. |  | 01: 0.029665; 07: 0.020094; 09: 0.026879; 11: 0.010488; 12: 0.018619; 13: 0.084004; 15: 0.037826; 17: 0.073598; 18: 0.010954
| Rel.Std.Dev. |  | 01: 2.799%; 07: 1.949%; 09: 2.788%; 11: 1.11%; 12: 1.596%; 13: 7.231%; 15: 3.627%; 17: 7.816%; 18: 1.204%
| PDM3 |  | 01: 3.437%; 07: 0.599%; 09: -5.925%; 11: -7.785%; 12: 13.85%; 13: 13.36%; 15: 1.778%; 17: -8.11%; 18: -11.2%
| Z-Score (Absolute) |  | 0.23 | NA | NA | NA | 0 | NA | 0.52 | 0.67 | 1.07 | 1.03 | NA | 0.09 | 0.7 | 0.95
| Analytical results for V in OREAS 600 (Certified Value 12.9 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  |  | 1
| 1 | 1 | 16.4733 | 17 | 12 | 12 | 13 | 12 | 18.3488 | 12 | 15 | 12.4075 | 13 | 15 | 12 | 17 | 10 | 13.464 | 9 | 11 | 12 |  |  |  | 1
| 1 | 2 | 16.4433 | 17 | 13 | 12 | 12 | 12 | 17.9308 | 13 | 14 | 12.2953 | 13 | 15 | 11 | 17 | 10 | 13.679 | 9 | 11 | 12 |  |  |  | 0
| 1 | 3 | 16.26 | 17 | 12 | 13 | 12 | 12 | 17.8673 | 11 | 14 | 12.2527 | 13 | 15 | 10 | 15 | 10 | 13.285 | 9 | 11 | 12 |  |  |  | 16
| 1 | 4 | 16.17 | 17 | 13 | 12 | 13 | 12 | 18.1776 | 12 | 14 | 12.98 | 13 | 15 | 13 | 17 | 10 | 13.468 | 9 | 11 | 14 |  |  |  | 12.9328
| 1 | 5 | 16.2233 | 16 | 12 | 13 | 12 | 12 | 18.6819 | 11 | 14 | 12.4449 | 13 | 15 | 11 | 16 | 10 | 13.861 | 10 | 11 | 12 |  |  |  | 116
| 1 | 6 | 16.56 | 17 | 12 | 12 | 13 | 12 | 18.2514 | 12 | 14 | 12.6726 | 12 | 16 | 12 | 17 | 10 | 13.803 | 10 | 11 | 12
| Mean |  | 16.355 | 16.8333 | 12.3333 | 12.3333 | 12.5 | 12 | 18.2096 | 11.8333 | 14.1667 | 12.5088 | 12.8333 | 15.1667 | 11.5 | 16.5 | 10 | 13.5933 | 9.3333 | 11 | 12.3333
| Median |  | 16.3517 | 17 | 12 | 12 | 12.5 | 12 | 18.2145 | 12 | 14 | 12.4262 | 13 | 15 | 11.5 | 17 | 10 | 13.5735 | 9 | 11 | 12
| Std Dev. |  | 0.15774 | 0.40825 | 0.5164 | 0.5164 | 0.54772 | 0 | 0.29668 | 0.75277 | 0.40825 | 0.27363 | 0.40825 | 0.40825 | 1.0488 | 0.83666 | 0 | 0.22381 | 0.5164 | 0 | 0.8165
| Rel.Std.Dev. |  | 0.9645% | 2.425% | 4.187% | 4.187% | 4.382% | 0% | 1.629% | 6.361% | 2.882% | 2.188% | 3.181% | 2.692% | 9.12% | 5.071% | 0% | 1.646% | 5.533% | 0% | 6.62%
| PDM3 |  | 26.46% | 30.16% | -4.635% | -4.635% | -3.347% | -7.213% | 40.8% | -8.501% | 9.541% | -3.278% | -0.7691% | 17.27% | -11.08% | 27.58% | -22.68% | 5.107% | -27.83% | -14.94% | -4.635%
| Z-Score (Absolute) |  | 2.38 | 2.67 | 0.1 | 0.1 | 0 | 0.31 | 3.52 | 0.41 | 1.03 | 0.01 | 0.21 | 1.64 | 0.62 | 2.47 | 1.54 | 0.67 | 1.95 | 0.93 | 0.31
| Analytical results for W in OREAS 600 (Indicative Value 0.46 ppm) | [23 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 25 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g |  |  |  |  | 2
| 1 | 1 | < 5 | 0.5 | < 10 | < 10 | < 10 | < 10 | 0.32242 | < 10 | < 10 | 0.58291 | < 2 | < 1 | < 10 | 0.8 | < 10 | 1.33 | < 2 | 0.2 |  |  |  |  | 1
| 1 | 2 | < 5 | 0.53 | < 10 | < 10 | < 10 | < 10 | 0.35282 | < 10 | < 10 | 0.58332 | < 2 | < 1 | < 10 | 0.6 | < 10 | 1.32 | < 2 | 0.2 |  |  |  |  | 9
| 1 | 3 | < 5 | 0.59 | < 10 | < 10 | < 10 | < 10 | 0.25712 | < 10 | < 10 | 0.5679 | < 2 | < 1 | < 10 | 0.6 | < 10 | 1.28 | < 2 | 0.3 |  |  |  |  | 16
| 1 | 4 | < 5 | 0.62 | < 10 | < 10 | < 10 | < 10 | 0.31625 | < 10 | < 10 | 0.53989 | < 2 | < 1 | < 10 | 0.5 | < 10 | 1.21 | < 2 | 0.3 |  |  |  |  | 0.45695
| 1 | 5 | < 5 | 0.57 | < 10 | < 10 | < 10 | < 10 | 0.32792 | < 10 | < 10 | 0.61824 | < 2 | < 1 | < 10 | 0.5 | < 10 | 1.16 | < 2 | 0.2 |  |  |  |  | 15
| 1 | 6 | < 5 | 0.59 | < 10 | < 10 | < 10 | < 10 | 0.40537 | < 10 | < 10 | 0.57449 | < 2 | < 1 | < 10 | 0.6 | < 10 | 1.16 | < 2 | 0.3
| Mean |  | 02: 0.56667; 09: 0.33032; 13: 0.57779; 17: 0.6; 19: 1.2433; 25: 0.25
| Median |  | 02: 0.58; 09: 0.32517; 13: 0.5787; 17: 0.6; 19: 1.245; 25: 0.25
| Std Dev. |  | 02: 0.044121; 09: 0.048505; 13: 0.025448; 17: 0.10954; 19: 0.077115; 25: 0.054772
| Rel.Std.Dev. |  | 02: 7.786%; 09: 14.68%; 13: 4.404%; 17: 18.26%; 19: 6.202%; 25: 21.91%
| PDM3 |  | 02: 24.01%; 09: -27.71%; 13: 26.44%; 17: 31.3%; 19: 172.1%; 25: -45.29%
| Z-Score (Absolute) |  | NA | 0.02 | NA | NA | NA | NA | 1.27 | NA | NA | 0.08 | NA | NA | NA | 0.02 | NA | 3.71 | NA | 1.71
| Analytical results for Y in OREAS 600 (Certified Value 5.98 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [9 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 40g | [9 empty] | 3
| 1 | 1 | 6.7171 | 6.06 | 6.0922 | 6 | 5.5984 | 6 | 6 | 5 | 3.98 | 4 | 5.984 | 6 | 6.47 | [9 empty] | 1
| 1 | 2 | 6.8286 | 6.25 | 5.938 | 6 | 5.794 | 6 | 6 | 5 | 4.13 | 4 | 6.063 | 6 | 6.27 | [9 empty] | 0
| 1 | 3 | 6.7337 | 6.06 | 5.8596 | 6 | 5.9407 | 6 | 6 | 5 | 3.47 | 4 | 6.101 | 6 | 6.61 | [9 empty] | 16
| 1 | 4 | 6.7595 | 6.04 | 6.0479 | 6 | 5.7191 | 6 | 6 | 5 | 4.23 | 4 | 5.962 | 6 | 5.91 | [9 empty] | 5.9785
| 1 | 5 | 6.7916 | 6.1 | 6.0537 | 5 | 6.0859 | 6 | 6 | 5 | 3.84 | 4 | 5.856 | 6 | 6.12 | [9 empty] | 117
| 1 | 6 | 6.7994 | 6.28 | 5.9404 | 6 | 6.017 | 6 | 6 | 5 | 4.11 | 4 | 5.895 | 6 | 5.83
| Mean |  | 6.7717 | 6.1317 | 5.9886 | 5.8333 | 5.8592 | 6 | 6 | 5 | 3.96 | 4 | 5.9768 | 6 | 6.2017
| Median |  | 6.7755 | 6.08 | 5.9941 | 6 | 5.8673 | 6 | 6 | 5 | 4.045 | 4 | 5.973 | 6 | 6.195
| Std Dev. |  | 0.042367 | 0.10553 | 0.089484 | 0.40825 | 0.18698 | 0 | 0 | 0 | 0.27539 | 0 | 0.094245 | 0 | 0.30766
| Rel.Std.Dev. |  | 0.6257% | 1.721% | 1.494% | 6.999% | 3.191% | 0% | 0% | 0% | 6.954% | 0% | 1.577% | 0% | 4.961%
| PDM3 |  | 13.27% | 2.563% | 0.1702% | -2.427% | -1.995% | 0.3605% | 0.3605% | -16.37% | -33.76% | -33.09% | -0.02704% | 0.3605% | 3.734%
| Z-Score (Absolute) |  | 3.69 | 0.67 | 0 | 0.73 | 0.61 | 0.05 | 0.05 | 4.66 | 9.56 | 9.37 | 0.06 | 0.05 | 1
| Analytical results for Yb in OREAS 600 (Certified Value 0.43 ppm) | [23 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 11: 25 |  |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*MS |  |  |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 11: 40g |  |  |  |  | 3
| 1 | 1 | 01: 0.4; 02: 0.44251; 05: 0.39001; 06: 0.45; 07: 0.42; 08: 0.4; 09: 0.406; 11: 0.42 |  |  |  |  | 1
| 1 | 2 | 01: 0.44; 02: 0.41281; 05: 0.37879; 06: 0.46; 07: 0.47; 08: 0.5; 09: 0.406; 11: 0.43 |  |  |  |  | 0
| 1 | 3 | 01: 0.43; 02: 0.42824; 05: 0.42903; 06: 0.45; 07: 0.43; 08: 0.4; 09: 0.413; 11: 0.42 |  |  |  |  | 16
| 1 | 4 | 01: 0.42; 02: 0.44337; 05: 0.36959; 06: 0.43; 07: 0.43; 08: 0.4; 09: 0.407; 11: 0.43 |  |  |  |  | 0.42656
| 1 | 5 | 01: 0.4; 02: 0.41242; 05: 0.4091; 06: 0.43; 07: 0.39; 08: 0.4; 09: 0.402; 11: 0.43 |  |  |  |  | 118
| 1 | 6 | 01: 0.42; 02: 0.44304; 05: 0.37073; 06: 0.42; 07: 0.44; 08: 0.5; 09: 0.413; 11: 0.39
| Mean |  | 01: 0.41833; 02: 0.4304; 05: 0.39121; 06: 0.44; 07: 0.43; 08: 0.43333; 09: 0.40783; 11: 0.42
| Median |  | 01: 0.42; 02: 0.43537; 05: 0.3844; 06: 0.44; 07: 0.43; 08: 0.4; 09: 0.4065; 11: 0.425
| Std Dev. |  | 01: 0.016021; 02: 0.014912; 05: 0.023632; 06: 0.015492; 07: 0.026077; 08: 0.05164; 09: 0.0043551; 11: 0.015492
| Rel.Std.Dev. |  | 01: 3.83%; 02: 3.465%; 05: 6.041%; 06: 3.521%; 07: 6.064%; 08: 11.92%; 09: 1.068%; 11: 3.689%
| PDM3 |  | 01: -1.928%; 02: 0.9005%; 05: -8.287%; 06: 3.152%; 07: 0.8072%; 08: 1.589%; 09: -4.389%; 11: -1.537%
| Z-Score (Absolute) |  | 01: 0.87; 02: 0.22; 05: 3.31; 06: 1.08; 07: 0.18; 08: 0.48; 09: 1.81; 11: 0.18
| Analytical results for Zn in OREAS 600 (Certified Value 598 ppm) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 0
| 1 | 1 | 630.486 | 595 | 597 | 635 | 640 | 617 | 615.5012 | 596 | 591 | 661 | 537.017 | 572.2 | 589 | 535 | 588 | 550 | 620.937 | 549 | 603 | 553 | 572 |  | 1
| 1 | 2 | 626.3325 | 632 | 629 | 646 | 639 | 624 | 610.7198 | 598 | 593 | 638 | 541.25 | 580.9 | 588 | 543 | 592 | 561 | 640.898 | 566 | 591 | 612 | 589 |  | 0
| 1 | 3 | 625.4227 | 618 | 639 | 634 | 654 | 612 | 604.4451 | 598 | 589 | 613 | 532.014 | 594.6 | 595 | 535 | 524 | 553 | 665.096 | 543 | 597 | 606 | 581 |  | 16
| 1 | 4 | 629.505 | 613 | 638 | 659 | 672 | 604 | 593.2764 | 522 | 583 | 641 | 540.573 | 584.2 | 588 | 539 | 594 | 553 | 648.905 | 561 | 606 | 601 | 584 |  | 597.7687
| 1 | 5 | 627.498 | 601 | 598 | 649 | 637 | 620 | 608.9765 | 602 | 582 | 637 | 533.788 | 599.7 | 594 | 552 | 569 | 556 | 656.699 | 560 | 596 | 572 | 567 |  | 119
| 1 | 6 | 627.057 | 601 | 633 | 656 | 654 | 604 | 616.2667 | 570 | 585 | 634 | 535.944 | 578.1 | 603 | 540 | 556 | 561 | 650.742 | 568 | 610 | 613 | 578
| Mean |  | 627.7169 | 610 | 622.3333 | 646.5 | 649.3333 | 613.5 | 608.1976 | 581 | 587.1667 | 637.3333 | 536.7643 | 584.95 | 592.8333 | 540.6667 | 570.5 | 555.6667 | 647.2128 | 557.8333 | 600.5 | 592.8333 | 578.5
| Median |  | 627.2775 | 607 | 631 | 647.5 | 647 | 614.5 | 609.8482 | 597 | 587 | 637.5 | 536.4805 | 582.55 | 591.5 | 539.5 | 578.5 | 554.5 | 649.8235 | 560.5 | 600 | 603.5 | 579.5
| Std Dev. |  | 1.925 | 13.7405 | 19.5721 | 10.4067 | 13.4412 | 8.3367 | 8.5105 | 31.132 | 4.4907 | 15.3449 | 3.6548 | 10.3647 | 5.8452 | 6.3456 | 27.1864 | 4.5461 | 15.2014 | 9.8268 | 7.064 | 24.6367 | 8.0187
| Rel.Std.Dev. |  | 0.3067% | 2.253% | 3.145% | 1.61% | 2.07% | 1.359% | 1.399% | 5.358% | 0.7648% | 2.408% | 0.6809% | 1.772% | 0.986% | 1.174% | 4.765% | 0.8181% | 2.349% | 1.762% | 1.176% | 4.156% | 1.386%
| PDM3 |  | 5.01% | 2.046% | 4.109% | 8.152% | 8.626% | 2.632% | 1.745% | -2.805% | -1.774% | 6.619% | -10.21% | -2.144% | -0.8256% | -9.553% | -4.562% | -7.043% | 8.271% | -6.681% | 0.4569% | -0.8256% | -3.223%
| Z-Score (Absolute) |  | 1.05 | 0.52 | 0.89 | 1.62 | 1.71 | 0.62 | 0.46 | 0 | 0.17 | 1.34 | 1.69 | 0.24 | 0 | 1.58 | 0.67 | 1.12 | 1.64 | 1.06 | 0.23 | 0 | 0.43
| Analytical results for Zr in OREAS 600 (Certified Value 10.2 ppm) | [23 empty] | Certified
Zr | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 12; 08: 13; 09: 14; 10: 15; 11: 16; 12: 18; 13: 19; 14: 25 |  | ppm
|  |  | 01: AR*OES; 02: AR*MS; 05: AR*OES/MS; 06: AR*OES; 07: AR*OES; 08: AR*MS; 09: AR*OES; 10: AR*OES/MS; 11: AR*OES; 12: AR*OES; 13: AR*OES/MS; 14: AR*MS |  | 1
|  |  | 01: 0.5g; 02: 40g; 05: 20g; 06: 0.15g; 07: 0.5g; 08: 01g; 09: 0.5g; 10: 01g; 11: 10g; 12: 10g; 13: 0.25g; 14: 40g |  | 1
| 1 | 1 | 01: 11.6375; 02: 10.7; 05: 8.9804; 06: 5; 07: 14; 08: 7.3238; 09: 10; 10: 11.22; 11: 10.7; 12: 11; 13: 7.343; 14: 10 |  | 1
| 1 | 2 | 01: 11.6525; 02: 10.4; 05: 8.5784; 06: 5; 07: 14; 08: 7.2182; 09: 10; 10: 10.38; 11: 10.3; 12: 12; 13: 6.45; 14: 10 |  | 0
| 1 | 3 | 01: 11.6575; 02: 10.9; 05: 9.1042; 06: 6; 07: 14; 08: 7.5048; 09: 10; 10: 10.1; 11: 10.3; 12: 12; 13: 6.537; 14: 10 |  | 16
| 1 | 4 | 01: 11.9175; 02: 11.1; 05: 9.0249; 06: 6; 07: 14; 08: 7.1085; 09: 10; 10: 11.05; 11: 10.3; 12: 12; 13: 6.543; 14: 10 |  | 10.2144
| 1 | 5 | 01: 11.93; 02: 10.6; 05: 9.8865; 06: 5; 07: 14; 08: 7.6787; 09: 10; 10: 10.74; 11: 10.2; 12: 11; 13: 6.939; 14: 10 |  | 120
| 1 | 6 | 01: 11.67; 02: 11.3; 05: 9.522; 06: 6; 07: 14; 08: 7.6447; 09: 9; 10: 10.38; 11: 10.3; 12: 12; 13: 6.749; 14: 10
| Mean |  | 01: 11.7442; 02: 10.8333; 05: 9.1828; 06: 5.5; 07: 14; 08: 7.4131; 09: 9.8333; 10: 10.645; 11: 10.35; 12: 11.6667; 13: 6.7602; 14: 10
| Median |  | 01: 11.6638; 02: 10.8; 05: 9.0646; 06: 5.5; 07: 14; 08: 7.4143; 09: 10; 10: 10.56; 11: 10.3; 12: 12; 13: 6.646; 14: 10
| Std Dev. |  | 01: 0.13955; 02: 0.33267; 05: 0.45776; 06: 0.54772; 07: 0; 08: 0.23295; 09: 0.40825; 10: 0.43385; 11: 0.17607; 12: 0.5164; 13: 0.33615; 14: 0
| Rel.Std.Dev. |  | 01: 1.188%; 02: 3.071%; 05: 4.985%; 06: 9.959%; 07: 0%; 08: 3.142%; 09: 4.152%; 10: 4.076%; 11: 1.701%; 12: 4.426%; 13: 4.972%; 14: 0%
| PDM3 |  | 01: 14.98%; 02: 6.059%; 05: -10.1%; 06: -46.15%; 07: 37.06%; 08: -27.42%; 09: -3.731%; 10: 4.215%; 11: 1.327%; 12: 14.22%; 13: -33.82%; 14: -2.099%
| Z-Score (Absolute) |  | 01: 0.87; 02: 0.38; 05: 0.52; 06: 2.52; 07: 2.1; 08: 1.48; 09: 0.17; 10: 0.27; 11: 0.08; 12: 0.83; 13: 1.84; 14: 0.08
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C103: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C121: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C157: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C175: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C193: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C211: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C229: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C247: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C283: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C301: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C319: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C337: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C355: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C373: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C409: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C427: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C445: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C463: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C481: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C499: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C535: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C553: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C571: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C589: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C661: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1183: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1201: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1219: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on D57: The INAA data is omitted from determination of the certified value for Au and is used solely for the calculation of Tolerance Limits and homogeneity evaluation of OREAS 600.

## Laser Ablation
| Analytical results for Ag in OREAS 600 (Indicative Value 20.9 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 21.7 | 1
| Mean |  | 20.85 | 16
| Median |  | 20.85 | 20.85
| Std Dev. |  | 1.2021 | 7
| Rel.Std.Dev. |  | 5.765%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 600 (Indicative Value 85 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84.6 | 1
| 1 | 2 | 85.2 | 2
| Mean |  | 84.9 | 16
| Median |  | 84.9 | 84.9
| Std Dev. |  | 0.42426 | 8
| Rel.Std.Dev. |  | 0.4997%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 600 (Indicative Value 1855 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1800 | 1
| 1 | 2 | 1910 | 3
| Mean |  | 1855 | 16
| Median |  | 1855 | 1855
| Std Dev. |  | 77.7817 | 9
| Rel.Std.Dev. |  | 4.193%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 600 (Indicative Value 0.8 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 1 | 4
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0.28284 | 10
| Rel.Std.Dev. |  | 35.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 600 (Indicative Value 6.23 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.02 | 1
| 1 | 2 | 6.44 | 5
| Mean |  | 6.23 | 16
| Median |  | 6.23 | 6.23
| Std Dev. |  | 0.29698 | 11
| Rel.Std.Dev. |  | 4.767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 600 (Indicative Value 2.95 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.1 | 1
| 1 | 2 | 2.8 | 6
| Mean |  | 2.95 | 16
| Median |  | 2.95 | 2.95
| Std Dev. |  | 0.21213 | 12
| Rel.Std.Dev. |  | 7.191%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 600 (Indicative Value 48.1 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 47 | 1
| 1 | 2 | 49.1 | 7
| Mean |  | 48.05 | 16
| Median |  | 48.05 | 48.05
| Std Dev. |  | 1.4849 | 13
| Rel.Std.Dev. |  | 3.09%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 600 (Indicative Value 7.1 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.8 | 1
| 1 | 2 | 7.4 | 8
| Mean |  | 7.1 | 16
| Median |  | 7.1 | 7.1
| Std Dev. |  | 0.42426 | 14
| Rel.Std.Dev. |  | 5.976%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 600 (Indicative Value 46 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 42 | 9
| Mean |  | 46 | 16
| Median |  | 46 | 46
| Std Dev. |  | 5.6569 | 15
| Rel.Std.Dev. |  | 12.3%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 600 (Indicative Value 9.34 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.26 | 1
| 1 | 2 | 9.41 | 10
| Mean |  | 9.335 | 16
| Median |  | 9.335 | 9.335
| Std Dev. |  | 0.10607 | 16
| Rel.Std.Dev. |  | 1.136%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 600 (Indicative Value 461 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 458 | 1
| 1 | 2 | 464 | 11
| Mean |  | 461 | 16
| Median |  | 461 | 461
| Std Dev. |  | 4.2426 | 17
| Rel.Std.Dev. |  | 0.9203%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 600 (Indicative Value 2.04 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.06 | 1
| 1 | 2 | 2.02 | 12
| Mean |  | 2.04 | 16
| Median |  | 2.04 | 2.04
| Std Dev. |  | 0.028284 | 18
| Rel.Std.Dev. |  | 1.386%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 600 (Indicative Value 1.03 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.96 | 1
| 1 | 2 | 1.09 | 13
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.091924 | 19
| Rel.Std.Dev. |  | 8.968%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 600 (Indicative Value 0.85 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.78 | 1
| 1 | 2 | 0.92 | 14
| Mean |  | 0.85 | 16
| Median |  | 0.85 | 0.85
| Std Dev. |  | 0.098995 | 20
| Rel.Std.Dev. |  | 11.65%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 600 (Indicative Value 18.4 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.9 | 1
| 1 | 2 | 18.8 | 15
| Mean |  | 18.35 | 16
| Median |  | 18.35 | 18.35
| Std Dev. |  | 0.6364 | 21
| Rel.Std.Dev. |  | 3.468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 600 (Indicative Value 2.91 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.89 | 1
| 1 | 2 | 2.93 | 16
| Mean |  | 2.91 | 16
| Median |  | 2.91 | 2.91
| Std Dev. |  | 0.028284 | 22
| Rel.Std.Dev. |  | 0.972%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 600 (Indicative Value 3.78 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.74 | 1
| 1 | 2 | 3.81 | 17
| Mean |  | 3.775 | 16
| Median |  | 3.775 | 3.775
| Std Dev. |  | 0.049497 | 23
| Rel.Std.Dev. |  | 1.311%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 600 (Indicative Value 0.36 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.37 | 1
| Mean |  | 0.355 | 16
| Median |  | 0.355 | 0.355
| Std Dev. |  | 0.021213 | 7
| Rel.Std.Dev. |  | 5.976%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 600 (Indicative Value 0.73 ppm) |  |  | Indicative
In | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.65 | 1
| 1 | 2 | 0.8 | 2
| Mean |  | 0.725 | 16
| Median |  | 0.725 | 0.725
| Std Dev. |  | 0.10607 | 8
| Rel.Std.Dev. |  | 14.63%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 600 (Indicative Value 27 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.2 | 1
| 1 | 2 | 27.7 | 3
| Mean |  | 26.95 | 16
| Median |  | 26.95 | 26.95
| Std Dev. |  | 1.0607 | 9
| Rel.Std.Dev. |  | 3.936%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 600 (Indicative Value 0.14 ppm) |  |  | Indicative
Lu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.12 | 1
| 1 | 2 | 0.15 | 4
| Mean |  | 0.135 | 16
| Median |  | 0.135 | 0.135
| Std Dev. |  | 0.021213 | 10
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 600 (Indicative Value 0.082 wt.%) |  |  | Indicative
Mn | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0844 | 1
| 1 | 2 | 0.0802 | 5
| Mean |  | 0.0823 | 16
| Median |  | 0.0823 | 0.0823
| Std Dev. |  | 0.0029698 | 11
| Rel.Std.Dev. |  | 3.609%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 600 (Indicative Value 2.4 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.2 | 6
| Mean |  | 2.4 | 16
| Median |  | 2.4 | 2.4
| Std Dev. |  | 0.28284 | 12
| Rel.Std.Dev. |  | 11.79%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 600 (Indicative Value 8.27 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.05 | 1
| 1 | 2 | 8.48 | 7
| Mean |  | 8.265 | 16
| Median |  | 8.265 | 8.265
| Std Dev. |  | 0.30406 | 13
| Rel.Std.Dev. |  | 3.679%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 600 (Indicative Value 20 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.5 | 1
| 1 | 2 | 19.5 | 8
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0.70711 | 14
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 600 (Indicative Value 17 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 16 | 1
| 1 | 2 | 18 | 9
| Mean |  | 17 | 16
| Median |  | 17 | 17
| Std Dev. |  | 1.4142 | 15
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 600 (Indicative Value 190 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 185 | 1
| 1 | 2 | 195 | 10
| Mean |  | 190 | 16
| Median |  | 190 | 190
| Std Dev. |  | 7.0711 | 16
| Rel.Std.Dev. |  | 3.722%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 600 (Indicative Value 5.59 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.33 | 1
| 1 | 2 | 5.84 | 11
| Mean |  | 5.585 | 16
| Median |  | 5.585 | 5.585
| Std Dev. |  | 0.36062 | 17
| Rel.Std.Dev. |  | 6.457%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 600 (Indicative Value 73 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72.1 | 1
| 1 | 2 | 73.3 | 12
| Mean |  | 72.7 | 16
| Median |  | 72.7 | 72.7
| Std Dev. |  | 0.84853 | 18
| Rel.Std.Dev. |  | 1.167%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 600 (Indicative Value 0.008 ppm) |  |  | Indicative
Re | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | < 0.01 | 13
| Mean |  | 0.01 | 16
| Median |  | 0.01 | 0.0075
| Std Dev. |  |  | 19
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 600 (Indicative Value 13.6 ppm) |  |  | Indicative
Sb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.6 | 1
| 1 | 2 | 13.5 | 14
| Mean |  | 13.55 | 16
| Median |  | 13.55 | 13.55
| Std Dev. |  | 0.070711 | 20
| Rel.Std.Dev. |  | 0.5219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 600 (Indicative Value 6 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.7 | 1
| 1 | 2 | 6.3 | 15
| Mean |  | 6 | 16
| Median |  | 6 | 6
| Std Dev. |  | 0.42426 | 21
| Rel.Std.Dev. |  | 7.071%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 600 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 16
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 22
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 600 (Indicative Value 3.71 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.74 | 1
| 1 | 2 | 3.67 | 17
| Mean |  | 3.705 | 16
| Median |  | 3.705 | 3.705
| Std Dev. |  | 0.049497 | 23
| Rel.Std.Dev. |  | 1.336%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 600 (Indicative Value 3 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3.2 | 1
| Mean |  | 3 | 16
| Median |  | 3 | 3
| Std Dev. |  | 0.28284 | 7
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 600 (Indicative Value 190 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 185 | 1
| 1 | 2 | 194 | 2
| Mean |  | 189.5 | 16
| Median |  | 189.5 | 189.5
| Std Dev. |  | 6.364 | 8
| Rel.Std.Dev. |  | 3.358%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 600 (Indicative Value 0.67 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.73 | 3
| Mean |  | 0.665 | 16
| Median |  | 0.665 | 0.665
| Std Dev. |  | 0.091924 | 9
| Rel.Std.Dev. |  | 13.82%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 600 (Indicative Value 0.35 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.33 | 4
| Mean |  | 0.345 | 16
| Median |  | 0.345 | 0.345
| Std Dev. |  | 0.021213 | 10
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 600 (Indicative Value 7.2 ppm) |  |  | Indicative
Te | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.6 | 1
| 1 | 2 | 6.8 | 5
| Mean |  | 7.2 | 16
| Median |  | 7.2 | 7.2
| Std Dev. |  | 0.56569 | 11
| Rel.Std.Dev. |  | 7.857%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 600 (Indicative Value 9.92 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.9 | 1
| 1 | 2 | 9.94 | 6
| Mean |  | 9.92 | 16
| Median |  | 9.92 | 9.92
| Std Dev. |  | 0.028284 | 12
| Rel.Std.Dev. |  | 0.2851%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 600 (Indicative Value 0.249 wt.%) |  |  | Indicative
Ti | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.243 | 1
| 1 | 2 | 0.255 | 7
| Mean |  | 0.249 | 16
| Median |  | 0.249 | 0.249
| Std Dev. |  | 0.0084853 | 13
| Rel.Std.Dev. |  | 3.408%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 600 (Indicative Value 1.1 ppm) |  |  | Indicative
Tl | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1 | 8
| Mean |  | 1.1 | 16
| Median |  | 1.1 | 1.1
| Std Dev. |  | 0.14142 | 14
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 600 (Indicative Value 0.15 ppm) |  |  | Indicative
Tm | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.17 | 1
| 1 | 2 | 0.13 | 9
| Mean |  | 0.15 | 16
| Median |  | 0.15 | 0.15
| Std Dev. |  | 0.028284 | 15
| Rel.Std.Dev. |  | 18.86%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 600 (Indicative Value 2.76 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.65 | 1
| 1 | 2 | 2.86 | 10
| Mean |  | 2.755 | 16
| Median |  | 2.755 | 2.755
| Std Dev. |  | 0.14849 | 16
| Rel.Std.Dev. |  | 5.39%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 600 (Indicative Value 47.1 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45.6 | 1
| 1 | 2 | 48.5 | 11
| Mean |  | 47.05 | 16
| Median |  | 47.05 | 47.05
| Std Dev. |  | 2.0506 | 17
| Rel.Std.Dev. |  | 4.358%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 600 (Indicative Value 3.68 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 3.55 | 12
| Mean |  | 3.675 | 16
| Median |  | 3.675 | 3.675
| Std Dev. |  | 0.17678 | 18
| Rel.Std.Dev. |  | 4.81%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 600 (Indicative Value 10.8 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.3 | 1
| 1 | 2 | 11.2 | 13
| Mean |  | 10.75 | 16
| Median |  | 10.75 | 10.75
| Std Dev. |  | 0.6364 | 19
| Rel.Std.Dev. |  | 5.92%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 600 (Indicative Value 0.97 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 0.93 | 14
| Mean |  | 0.965 | 16
| Median |  | 0.965 | 0.965
| Std Dev. |  | 0.049497 | 20
| Rel.Std.Dev. |  | 5.129%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 600 (Indicative Value 580 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 580 | 1
| 1 | 2 | 580 | 15
| Mean |  | 580 | 16
| Median |  | 580 | 580
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 600 (Indicative Value 134 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 131 | 1
| 1 | 2 | 137 | 16
| Mean |  | 134 | 16
| Median |  | 134 | 134
| Std Dev. |  | 4.2426 | 22
| Rel.Std.Dev. |  | 3.166%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 600 (Indicative Value 13.39 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 13.4 | 1
| 1 | 2 | 13.38 | 19
| Mean |  | 13.39 | 16
| Median |  | 13.39 | 13.39
| Std Dev. |  | 0.014142 | 25
| Rel.Std.Dev. |  | 0.1056%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 600 (Indicative Value 100 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 2
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 600 (Indicative Value 1865 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1870 | 1
| 1 | 2 | 1860 | 3
| Mean |  | 1865 | 16
| Median |  | 1865 | 1865
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 0.3791%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 600 (Indicative Value 2.7 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.7 | 1
| 1 | 2 | 2.69 | 22
| Mean |  | 2.695 | 16
| Median |  | 2.695 | 2.695
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.2624%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 600 (Indicative Value 10 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 8
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 600 (Indicative Value 40 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 40 | 1
| 1 | 2 | 40 | 9
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 600 (Indicative Value 475 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 480 | 1
| 1 | 2 | 470 | 11
| Mean |  | 475 | 16
| Median |  | 475 | 475
| Std Dev. |  | 7.0711 | 31
| Rel.Std.Dev. |  | 1.489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 600 (Indicative Value 3.44 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.43 | 1
| 1 | 2 | 3.44 | 19
| Mean |  | 3.435 | 16
| Median |  | 3.435 | 3.435
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.2059%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 600 (Indicative Value 2.24 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.24 | 1
| 1 | 2 | 2.23 | 20
| Mean |  | 2.235 | 16
| Median |  | 2.235 | 2.235
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 0.3164%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 600 (Indicative Value 1.36 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.36 | 1
| 1 | 2 | 1.35 | 21
| Mean |  | 1.355 | 16
| Median |  | 1.355 | 1.355
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 0.5219%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 600 (Indicative Value 0.105 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.11 | 22
| Mean |  | 0.105 | 16
| Median |  | 0.105 | 0.105
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 600 (Indicative Value 0.835 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.83 | 23
| Mean |  | 0.835 | 16
| Median |  | 0.835 | 0.835
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.8468%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 600 (Indicative Value < 10 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 9
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 600 (Indicative Value 0.143 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.143 | 1
| 1 | 2 | 0.143 | 25
| Mean |  | 0.143 | 16
| Median |  | 0.143 | 0.143
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 600 (Indicative Value 190 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 190 | 1
| 1 | 2 | 190 | 10
| Mean |  | 190 | 16
| Median |  | 190 | 190
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 600 (Indicative Value 68.05 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 68.1 | 1
| 1 | 2 | 67.99 | 20
| Mean |  | 68.045 | 16
| Median |  | 68.045 | 68.045
| Std Dev. |  | 0.077782 | 26
| Rel.Std.Dev. |  | 0.1143%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 600 (Indicative Value 12.5 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | < 10 | 1
| Mean |  | 20 | 16
| Median |  | 20 | 12.5
| Std Dev. |  |  | 27
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 600 (Indicative Value 4.11 wt.%) |  |  | Indicative
SO3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.101 | 1
| 1 | 2 | 4.116 | 22
| Mean |  | 4.1085 | 16
| Median |  | 4.1085 | 4.1085
| Std Dev. |  | 0.010607 | 28
| Rel.Std.Dev. |  | 0.2582%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 600 (Indicative Value 0.417 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.418 | 1
| 1 | 2 | 0.416 | 23
| Mean |  | 0.417 | 16
| Median |  | 0.417 | 0.417
| Std Dev. |  | 0.0014142 | 29
| Rel.Std.Dev. |  | 0.3391%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 600 (Indicative Value < 10 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 10
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 600 (Indicative Value 585 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 590 | 1
| 1 | 2 | 580 | 15
| Mean |  | 585 | 16
| Median |  | 585 | 585
| Std Dev. |  | 7.0711 | 31
| Rel.Std.Dev. |  | 1.209%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 600 (Indicative Value 4.52 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | at 1000°C | 2
2
| 1 | 1 | 4.54 | 1
| 1 | 2 | 4.5 | 27
| Mean |  | 4.52 | 16
| Median |  | 4.52 | 4.52
| Std Dev. |  | 0.028284 | 33
| Rel.Std.Dev. |  | 0.6258%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 600 (Certified Value 24.8 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | AD*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 25.4962 | 23.46 | 24.6 | 25.1 | 24.8 | 23.3 | 23.5862 | 25.2 | 26.75 | 24.5766 | 24.6 | 24.56 | 27 | > 10 | 24 | 27.177 | 26 | 18 | 22.4 | 25 | 20.5 | 1
| 1 | 2 | 25.9 | 23.73 | 24.4 | 25.8 | 24 | 22.3 | 23.8843 | 25.2 | 24.79 | 25.9135 | 24.3 | 24.85 | 26 | > 10 | 24 | 26.272 | 26 | 21 | 23.6 | 30 | 20.5 | 1
| 1 | 3 | 25.91 | 23.41 | 24.7 | 24.5 | 24.1 | 23.7 | 24.0551 | 25.5 | 24.33 | 23.9811 | 24.4 | 23.84 | 25 | > 10 | 24 | 25.448 | 26 | 21 | 23.2 | 25 | 21.5 | 16
| 1 | 4 | 25.7013 | 23.37 | 25.2 | 24.7 | 24.3 | 24 | 24.3152 | 25.2 | 26.41 | 24.0982 | 24.1 | 24.56 | 25 | > 10 | 24 | 26.571 | 26 | 22 | 23.1 | 25 | 20 | 24.7513
| 1 | 5 | 25.0071 | 23.4 | 23.9 | 24.9 | 24.8 | 23.3 | 24.1955 | 25.8 | 27.06 | 24.1703 | 24.2 | 24.66 | 26 | > 10 | 25 | 26.087 | 25 | 21 | 24 | 25 | 20.5 | 9
| 1 | 6 | 25.198 | 23.47 | 25.2 | 25.2 | 24.6 | 24.2 | 23.9373 | 25.2 | 24.11 | 26.4046 | 24.3 | 24.08 | 26 | > 10 | 24 | 28.969 | 26 | 22 | 23.2 | 30 | 20.5
| Mean |  | 25.5354 | 23.4733 | 24.6667 | 25.0333 | 24.4333 | 23.4667 | 23.9956 | 25.35 | 25.575 | 24.8574 | 24.3167 | 24.425 | 25.8333 |  | 24.1667 | 26.754 | 25.8333 | 20.8333 | 23.25 | 26.6667 | 20.5833
| Median |  | 25.5987 | 23.435 | 24.65 | 25 | 24.45 | 23.5 | 23.9962 | 25.2 | 25.6 | 24.3735 | 24.3 | 24.56 | 26 |  | 24 | 26.4215 | 26 | 21 | 23.2 | 25 | 20.5
| Std Dev. |  | 0.37288 | 0.13125 | 0.49666 | 0.45461 | 0.35024 | 0.67725 | 0.25652 | 0.251 | 1.3111 | 1.0396 | 0.17224 | 0.38303 | 0.75277 |  | 0.40825 | 1.2249 | 0.40825 | 1.472 | 0.53572 | 2.582 | 0.4916
| Rel.Std.Dev. |  | 1.46% | 0.5591% | 2.013% | 1.816% | 1.433% | 2.886% | 1.069% | 0.9901% | 5.127% | 4.182% | 0.7083% | 1.568% | 2.914% |  | 1.689% | 4.578% | 1.58% | 7.065% | 2.304% | 9.682% | 2.388%
| PDM3 |  | 3.168% | -5.163% | -0.3418% | 1.14% | -1.285% | -5.19% | -3.053% | 2.419% | 3.328% | 0.4288% | -1.756% | -1.318% | 4.372% |  | -2.362% | 8.091% | 4.372% | -15.83% | -6.065% | 7.739% | -16.84%
| Z-Score (Absolute) |  | 0.74 | 0.81 | 0.09 | 0.37 | 0.09 | 0.82 | 0.42 | 0.6 | 0.77 | 0.23 | 0.18 | 0.09 | 0.97 | NA | 0.29 | 1.33 | 0.97 | 2.38 | 0.98 | 0.34 | 3
| Analytical results for Al in OREAS 600 (Certified Value 6.78 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 7.205 | 6.8515 | 6.6 | 6.58 | 6.8 | 6.55 | 6.6994 | 6.57 | 6.99 | 4.05 | 6.9978 | 6.9402 | 7.13 | 4.48 | 6.69 | 6.41 | 7.024 | 5.7 | 6.82 | 6.8 |  | 1
| 1 | 2 | 7.03 | 6.8731 | 6.63 | 6.73 | 6.54 | 6.23 | 6.7723 | 6.59 | 7 | 3.36 | 7.0759 | 6.9336 | 7.35 | 4.3 | 6.67 | 6.78 | 7.048 | 6.31 | 6.69 | 6.84 |  | 0
| 1 | 3 | 6.865 | 6.7817 | 6.66 | 6.44 | 6.4 | 6.57 | 6.9219 | 6.22 | 7.19 | 5.85 | 7.0041 | 6.7604 | 7.25 | 5.16 | 6.91 | 6.64 | 6.951 | 6.2 | 6.67 | 6.77 |  | 16
| 1 | 4 | 6.93 | 6.8944 | 6.86 | 6.49 | 6.68 | 6.65 | 6.8604 | 6.67 | 6.67 | 6.85 | 7.062 | 6.7942 | 7.23 | 4.54 | 6.94 | 6.75 | 6.773 | 6.23 | 6.79 | 6.89 |  | 6.7802
| 1 | 5 | 6.985 | 6.9113 | 6.55 | 6.39 | 6.61 | 6.51 | 6.8615 | 6.14 | 7.09 | 7.25 | 7.0484 | 6.8597 | 7.26 | 5.37 | 6.92 | 6.83 | 6.942 | 6.15 | 6.6 | 6.86 |  | 10
| 1 | 6 | 7.15 | 6.6545 | 6.85 | 6.62 | 6.46 | 6.68 | 6.7731 | 6.56 | 7.05 | 6.8 | 7.1204 | 6.7734 | 7.24 | 5.99 | 6.75 | 6.71 | 7.542 | 6.64 | 6.8 | 6.78
| Mean |  | 7.0275 | 6.8277 | 6.6917 | 6.5417 | 6.5817 | 6.5317 | 6.8148 | 6.4583 | 6.9983 | 5.6933 | 7.0514 | 6.8436 | 7.2433 | 4.9733 | 6.8133 | 6.6867 | 7.0467 | 6.205 | 6.7283 | 6.8233
| Median |  | 7.0075 | 6.8623 | 6.645 | 6.535 | 6.575 | 6.56 | 6.8167 | 6.565 | 7.025 | 6.325 | 7.0552 | 6.827 | 7.245 | 4.85 | 6.83 | 6.73 | 6.9875 | 6.215 | 6.74 | 6.82
| Std Dev. |  | 0.1298 | 0.096104 | 0.13167 | 0.12576 | 0.14675 | 0.1608 | 0.080741 | 0.22049 | 0.17646 | 1.6221 | 0.046039 | 0.079977 | 0.070333 | 0.64967 | 0.12372 | 0.15002 | 0.2611 | 0.30297 | 0.087958 | 0.04761
| Rel.Std.Dev. |  | 1.847% | 1.408% | 1.968% | 1.923% | 2.23% | 2.462% | 1.185% | 3.414% | 2.521% | 28.49% | 0.6529% | 1.169% | 0.971% | 13.06% | 1.816% | 2.244% | 3.705% | 4.883% | 1.307% | 0.6977%
| PDM3 |  | 3.648% | 0.7018% | -1.305% | -3.518% | -2.928% | -3.665% | 0.5102% | -4.747% | 3.218% | -16.03% | 4.001% | 0.9354% | 6.831% | -26.65% | 0.4892% | -1.379% | 3.931% | -8.483% | -0.7644% | 0.6367%
| Z-Score (Absolute) |  | 0.94 | 0.21 | 0.29 | 0.84 | 0.69 | 0.66 | 0.16 | 1.15 | 1.07 | 3.95 | 1.03 | 0.27 | 1.73 | 6.59 | 0.16 | 0.31 | 0.65 | 1.7 | 0.16 | 0.19
| Analytical results for As in OREAS 600 (Certified Value 89 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  | 0
1
| 1 | 1 | 93 | 90.7 | 90 | 85 | 88 | 89 | 93.2646 | 90 | 85 | 9.8 | 123 | 106 | 78 | 99 | 113 | 85.664 | 77 | 73 | 82 | 70 |  | 1
| 1 | 2 | 93 | 90.1 | 93 | 106 | 89 | 83 | 93.6203 | 90 | 82 | 9 | 124 | 101 | 75 | 95 | 92 | 85.749 | 88 | 65 | 79 | 74 |  | 0
| 1 | 3 | 92 | 91.9 | 93 | 87 | 85 | 89 | 93.8167 | 90 | 88 | 12.1 | 97 | 98 | 79 | 102 | 94 | 82.346 | 87 | 61 | 86 | 74 |  | 16
| 1 | 4 | 89 | 92.9 | 91 | 90 | 86 | 90 | 94.6604 | 90 | 81 | 15.6 | 94 | 97 | 74 | 104 | 100 | 82.361 | 96 | 70 | 88 | 76 |  | 88.8006
| 1 | 5 | 88 | 91.4 | 86 | 86 | 94 | 86 | 94.3227 | 90 | 87 | 16.1 | 96 | 98 | 79 | 106 | 100 | 77.668 | 94 | 65 | 87 | 75 |  | 11
| 1 | 6 | 89 | 91.3 | 96 | 87 | 89 | 93 | 94.2805 | 90 | 85 | 15.1 | 126 | 99 | 78 | 98 | 102 | 78.412 | 90 | 61 | 81 | 76
| Mean |  | 90.6667 | 91.3833 | 91.5 | 90.1667 | 88.5 | 88.3333 | 93.9942 | 90 | 84.6667 | 12.95 | 110 | 99.8333 | 77.1667 | 100.6667 | 100.1667 | 82.0333 | 88.6667 | 65.8333 | 83.8333 | 74.1667
| Median |  | 90.5 | 91.35 | 92 | 87 | 88.5 | 89 | 94.0486 | 90 | 85 | 13.6 | 110 | 98.5 | 78 | 100.5 | 100 | 82.3535 | 89 | 65 | 84 | 74.5
| Std Dev. |  | 2.2509 | 0.96833 | 3.3912 | 7.9352 | 3.1464 | 3.4448 | 0.51718 | 0 | 2.7325 | 3.0924 | 15.7607 | 3.3116 | 2.137 | 4.0825 | 7.3869 | 3.4457 | 6.6833 | 4.8339 | 3.656 | 2.2286
| Rel.Std.Dev. |  | 2.483% | 1.06% | 3.706% | 8.801% | 3.555% | 3.9% | 0.5502% | 0% | 3.227% | 23.88% | 14.33% | 3.317% | 2.769% | 4.055% | 7.375% | 4.2% | 7.538% | 7.343% | 4.361% | 3.005%
| PDM3 |  | 2.101% | 2.908% | 3.04% | 1.538% | -0.3386% | -0.5262% | 5.849% | 1.351% | -4.655% | -85.42% | 23.87% | 12.42% | -13.1% | 13.36% | 12.8% | -7.621% | -0.1509% | -25.86% | -5.594% | -16.48%
| Z-Score (Absolute) |  | 0.28 | 0.37 | 0.39 | 0.21 | 0.01 | 0.03 | 0.72 | 0.19 | 0.52 | 10.04 | 2.84 | 1.33 | 1.52 | 1.6 | 1.2 | 0.87 | 0.01 | 3.02 | 0.63 | 1.8
| Analytical results for B in OREAS 600 (Indicative Value < 20 ppm) | [22 empty] | Indicative
B | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 25 |  | ppm
|  |  | 01: PF*OES |  | 1
1
| 1 | 1 | 01: < 20 |  | 1
| 1 | 2 | 01: < 20 |  | 3
| 1 | 3 | 01: < 20 |  | 16
| 1 | 4 | 01: < 20 |  | < 20
| 1 | 5 | 01: < 20 |  | 9
| 1 | 6 | 01: < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for Ba in OREAS 600 (Indicative Value 1845 ppm) | [22 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 23 | 25 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*MS | [5 empty] | 0
0
| 1 | 1 | 1835.7 | 1360 | 200 | 1630 | 1350 | 1653.4665 | 869 | 775 | 1757.16 | 1903 | 821 | 478 | 1867 | 1086 | 301 | 2010 | [5 empty] | 1
| 1 | 2 | 1826 | 1350 | 210 | 1560 | 1260 | 1638.462 | 904 | 717 | 1771.263 | 1877 | 1164 | 488 | 1914 | 977 | 197 | 1970 | [5 empty] | 4
| 1 | 3 | 1870.1 | 1330 | 210 | 1580 | 1380 | 1628.1615 | 888 | 978 | 1749.195 | 1844 | 905 | 464 | 1883 | 1232 | 192 | 1820 | [5 empty] | 16
| 1 | 4 | 1860.8 | 1310 | 210 | 1550 | 1390 | 1626.597 | 1040 | 1161 | 1785.456 | 1840 | 976 | 488 | 1843 | 1438 | 169 | 1760 | [5 empty] | 1844.6411
| 1 | 5 | 1829.1 | 1340 | 190 | 470 | 1370 | 1661.8035 | 1157 | 1459 | 1757.691 | 1846 | 1546 | 461 | 1860 | 1356 | 145 | 1750 | [5 empty] | 10
| 1 | 6 | 1835.6 | 1450 | 210 | 490 | 1320 | 1658.1337 | 944 | 1225 | 1811.169 | 1894 | 1144 | 458 | 1919 | 1497 | 278 | 1850
| Mean |  | 1842.8833 | 1356.6667 | 205 | 1213.3333 | 1345 | 1644.4374 | 967 | 1052.5 | 1771.989 | 1867.3333 | 1092.6667 | 472.8333 | 1881 | 1264.3333 | 213.6667 | 1860
| Median |  | 1835.65 | 1345 | 210 | 1555 | 1360 | 1645.9642 | 924 | 1069.5 | 1764.477 | 1861.5 | 1060 | 471 | 1875 | 1294 | 194.5 | 1835
| Std Dev. |  | 18.1187 | 48.8535 | 8.3666 | 568.7413 | 48.4768 | 15.4246 | 111.1503 | 283.5093 | 23.0688 | 27.6526 | 259.1784 | 13.6002 | 30.3908 | 204.0124 | 62.0054 | 108.074
| Rel.Std.Dev. |  | 0.9832% | 3.601% | 4.081% | 46.87% | 3.604% | 0.938% | 11.49% | 26.94% | 1.302% | 1.481% | 23.72% | 2.876% | 1.616% | 16.14% | 29.02% | 5.81%
| PDM3 |  | -0.09529% | -26.45% | -88.89% | -34.22% | -27.09% | -10.85% | -47.58% | -42.94% | -3.939% | 1.23% | -40.77% | -74.37% | 1.971% | -31.46% | -88.42% | 0.8326%
| Z-Score (Absolute) |  | 0.9 | 0.06 | 1.81 | 0.15 | 0.07 | 0.57 | 0.55 | 0.41 | 0.78 | 0.94 | 0.35 | 1.37 | 0.96 | 0.06 | 1.8 | 0.93
| Analytical results for Be in OREAS 600 (Certified Value 1.69 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | 1.7891 | 1.73 | 1.5 | 1.6 | 1.6 | 1.6 | 1.5897 | 1.6 | 2 | 1.7263 | 1.5 | 2.03 | 1.5 | 3.1 | < 3 | 1.499 | 2 | 1 | 1.5 | 1.5 |  | 1
| 1 | 2 | 1.79 | 1.79 | 1.6 | 1.7 | 1.6 | 1.5 | 1.6879 | 1.7 | 1.45 | 1.6998 | 1.5 | 1.99 | 1.5 | 3.3 | < 3 | 1.781 | 3 | 2 | 2 | 1.5 |  | 0
| 1 | 3 | 1.7415 | 1.8 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6149 | 1.7 | 2.34 | 1.6857 | 1.4 | 1.95 | 1.5 | 3 | < 3 | 1.515 | 2 | 2 | 2 | 1.5 |  | 16
| 1 | 4 | 1.768 | 1.77 | 1.6 | 1.6 | 1.6 | 1.6 | 1.5662 | 2 | 1.93 | 1.695 | 1.6 | 2.03 | 1.5 | 3.1 | < 3 | 1.551 | 2 | 2 | 2 | 1.5 |  | 1.6914
| 1 | 5 | 1.7066 | 1.73 | 1.5 | 1.6 | 1.6 | 1.6 | 1.7141 | 1.7 | 1.79 | 1.6876 | 1.4 | 2.08 | 1.6 | 3 | < 3 | 1.786 | 2 | 2 | 2 | 1.5 |  | 12
| 1 | 6 | 1.7921 | 1.72 | 1.6 | 1.7 | 1.5 | 1.6 | 1.6697 | 1.7 | 1.81 | 1.6776 | 1.4 | 1.95 | 1.6 | 3.1 | < 3 | 1.561 | 2 | 2 | 2 | 1.6
| Mean |  | 1.7646 | 1.7567 | 1.5667 | 1.6333 | 1.5833 | 1.5833 | 1.6404 | 1.7333 | 1.8867 | 1.6953 | 1.4667 | 2.005 | 1.5333 | 3.1 |  | 1.6155 | 2.1667 | 1.8333 | 1.9167 | 1.5167
| Median |  | 1.7786 | 1.75 | 1.6 | 1.6 | 1.6 | 1.6 | 1.6423 | 1.7 | 1.87 | 1.6913 | 1.45 | 2.01 | 1.5 | 3.1 |  | 1.556 | 2 | 2 | 2 | 1.5
| Std Dev. |  | 0.034383 | 0.034448 | 0.05164 | 0.05164 | 0.040825 | 0.040825 | 0.0588 | 0.13663 | 0.29194 | 0.017015 | 0.08165 | 0.051284 | 0.05164 | 0.10954 |  | 0.13211 | 0.40825 | 0.40825 | 0.20412 | 0.040825
| Rel.Std.Dev. |  | 1.949% | 1.961% | 3.296% | 3.162% | 2.578% | 2.578% | 3.584% | 7.882% | 15.47% | 1.004% | 5.567% | 2.558% | 3.368% | 3.534% |  | 8.178% | 18.84% | 22.27% | 10.65% | 2.692%
| PDM3 |  | 4.326% | 3.86% | -7.373% | -3.432% | -6.388% | -6.388% | -3.013% | 2.48% | 11.55% | 0.233% | -13.29% | 18.54% | -9.344% | 83.28% |  | -4.486% | 28.1% | 8.393% | 13.32% | -10.33%
| Z-Score (Absolute) |  | 0.5 | 0.46 | 0.67 | 0.28 | 0.58 | 0.58 | 0.24 | 0 | 0.69 | 0.09 | 1.27 | 1.93 | 0.87 | 8.45 | NA | 0.38 | 2.9 | 1.9 | 1.41 | 0.97
| Analytical results for Bi in OREAS 600 (Certified Value 6.39 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 2
2
| 1 | 1 | < 5 | 6.69 | 8 | 5 | 6 | 6 | 6.3748 | < 10 | 5 | 8.27 | 6.0031 | 7.7 | 7.09 | < 30 | 6.89 | < 20 | 6.361 | 3 | < 5 | 5.8 | 6.3 | 1
| 1 | 2 | < 5 | 6.71 | 7 | 5 | 6 | 6 | 6.5608 | < 10 | 6 | 5.82 | 6.3279 | 7.15 | 7.63 | < 30 | 6.47 | < 20 | 6.576 | 3 | < 5 | 6 | 6.9 | 0
| 1 | 3 | < 5 | 6.57 | 7 | 6 | 5 | 6 | 6.4172 | 20 | 6 | 7.29 | 6.6654 | 6.51 | 7.6 | < 30 | 6.46 | < 20 | 6.221 | 2 | < 5 | 5.8 | 6.9 | 16
| 1 | 4 | < 5 | 6.91 | 9 | 6 | 7 | 4 | 6.3246 | < 10 | 6 | 7 | 6.132 | 6.04 | 7.72 | < 30 | 6.79 | < 20 | 6.285 | 3 | < 5 | 5.8 | 6.5 | 6.3898
| 1 | 5 | < 5 | 6.9 | 8 | 5 | 11 | 6 | 6.488 | 20 | 5 | 6.69 | 6.1092 | 6.3 | 7.24 | < 30 | 6.94 | < 20 | 6.377 | 3 | < 5 | 5.9 | 6.6 | 13
| 1 | 6 | < 5 | 6.15 | 5 | 6 | 11 | 10 | 6.3859 | < 10 | 5 | 6.59 | 6.0379 | 7.35 | 7.17 | < 30 | 6.64 | < 20 | 6.235 | 2 | < 5 | 6.3 | 6.5
| Mean |  |  | 6.655 | 7.3333 | 5.5 | 7.6667 | 6.3333 | 6.4252 | 20 | 5.5 | 6.9433 | 6.2126 | 6.8417 | 7.4083 |  | 6.6983 |  | 6.3425 | 2.6667 |  | 5.9333 | 6.6167
| Median |  |  | 6.7 | 7.5 | 5.5 | 6.5 | 6 | 6.4015 | 20 | 5.5 | 6.845 | 6.1206 | 6.83 | 7.42 |  | 6.715 |  | 6.323 | 3 |  | 5.85 | 6.55
| Std Dev. |  |  | 0.27984 | 1.3663 | 0.54772 | 2.6583 | 1.9664 | 0.085553 | 0 | 0.54772 | 0.81662 | 0.24897 | 0.65365 | 0.27184 |  | 0.20779 |  | 0.1309 | 0.5164 |  | 0.19664 | 0.24014
| Rel.Std.Dev. |  |  | 4.205% | 18.63% | 9.959% | 34.67% | 31.05% | 1.332% | 0% | 9.959% | 11.76% | 4.008% | 9.554% | 3.669% |  | 3.102% |  | 2.064% | 19.36% |  | 3.314% | 3.629%
| PDM3 |  |  | 4.151% | 14.77% | -13.92% | 19.98% | -0.883% | 0.5547% | 213% | -13.92% | 8.663% | -2.773% | 7.072% | 15.94% |  | 4.829% |  | -0.7396% | -58.27% |  | -7.143% | 3.551%
| Z-Score (Absolute) |  | NA | 0.3 | 1.18 | 1.2 | 0.55 | 1.07 | 0 | 17.67 | 1.2 | 0.67 | 0.39 | 0.54 | 1.28 | NA | 0.36 | NA | 0.11 | 4.89 | NA | 0.74 | 0.25
| Analytical results for Ca in OREAS 600 (Certified Value 1.88 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.055 | 1.9078 | 1.91 | 1.94 | 1.91 | 1.85 | 1.8832 | 1.78 | 1.89 | 1.95 | 1.7795 | 1.9342 | 1.91 | 1.48 | 1.8 | 1.74 | 1.914 | 2.28 | 1.66 | 1.98 | 1.92 | 1
| 1 | 2 | 2 | 1.8859 | 1.92 | 1.98 | 1.84 | 1.78 | 1.8713 | 1.79 | 1.92 | 1.73 | 1.7993 | 1.9003 | 1.91 | 1.45 | 1.79 | 1.84 | 1.963 | 2.53 | 1.76 | 1.97 | 1.9 | 0
| 1 | 3 | 1.92 | 1.8633 | 1.93 | 1.9 | 1.85 | 1.87 | 1.8404 | 1.73 | 1.98 | 1.87 | 1.7888 | 1.9255 | 1.91 | 1.54 | 1.86 | 1.8 | 1.859 | 2.54 | 1.71 | 1.98 | 1.88 | 16
| 1 | 4 | 1.995 | 1.9027 | 1.98 | 1.89 | 1.85 | 1.91 | 1.8669 | 1.83 | 1.9 | 1.93 | 1.7914 | 1.8715 | 1.89 | 1.52 | 1.86 | 1.81 | 1.88 | 2.54 | 1.75 | 1.99 | 1.92 | 1.8781
| 1 | 5 | 1.975 | 1.8692 | 1.88 | 1.9 | 1.91 | 1.86 | 1.8612 | 1.79 | 1.98 | 1.89 | 1.7878 | 1.9052 | 1.92 | 1.64 | 1.86 | 1.83 | 1.92 | 2.61 | 1.73 | 1.99 | 1.9 | 14
| 1 | 6 | 1.97 | 1.8695 | 1.95 | 1.96 | 1.88 | 1.92 | 1.9039 | 1.83 | 1.95 | 1.85 | 1.7952 | 1.8554 | 1.89 | 1.65 | 1.82 | 1.79 | 1.844 | 2.6 | 1.79 | 2.01 | 1.92
| Mean |  | 1.9858 | 1.8831 | 1.9283 | 1.9283 | 1.8733 | 1.865 | 1.8711 | 1.7917 | 1.9367 | 1.87 | 1.7903 | 1.8987 | 1.905 | 1.5467 | 1.8317 | 1.8017 | 1.8967 | 2.5167 | 1.7333 | 1.9867 | 1.9067
| Median |  | 1.985 | 1.8777 | 1.925 | 1.92 | 1.865 | 1.865 | 1.8691 | 1.79 | 1.935 | 1.88 | 1.7901 | 1.9027 | 1.91 | 1.53 | 1.84 | 1.805 | 1.897 | 2.54 | 1.74 | 1.985 | 1.91
| Std Dev. |  | 0.044206 | 0.018827 | 0.034303 | 0.037103 | 0.031411 | 0.0501 | 0.021365 | 0.037103 | 0.039328 | 0.077974 | 0.0067971 | 0.030457 | 0.012247 | 0.082381 | 0.032506 | 0.035449 | 0.04408 | 0.12078 | 0.045019 | 0.013663 | 0.01633
| Rel.Std.Dev. |  | 2.226% | 0.9998% | 1.779% | 1.924% | 1.677% | 2.686% | 1.142% | 2.071% | 2.031% | 4.17% | 0.3797% | 1.604% | 0.6429% | 5.326% | 1.775% | 1.968% | 2.324% | 4.799% | 2.597% | 0.6877% | 0.8565%
| PDM3 |  | 5.736% | 0.264% | 2.674% | 2.674% | -0.2542% | -0.6979% | -0.3708% | -4.603% | 3.118% | -0.4317% | -4.674% | 1.096% | 1.432% | -17.65% | -2.473% | -4.07% | 0.9882% | 34% | -7.709% | 5.78% | 1.521%
| Z-Score (Absolute) |  | 1.53 | 0 | 0.67 | 0.67 | 0.14 | 0.27 | 0.18 | 1.36 | 0.8 | 0.19 | 1.38 | 0.23 | 0.33 | 5.01 | 0.77 | 1.21 | 0.2 | 10.14 | 2.23 | 1.54 | 0.35
| Analytical results for Cd in OREAS 600 (Certified Value 3.37 ppm) | [22 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 3.0835 | 3.36 | 3.4 | 4.3 | 3.5 | 3.3 | 3.4375 | 5 | 3 | 3.94 | 3.4981 | 3.95 | 4.17 | 3 | 3.81 | 4 | 3.343 | < 4 | 3.2 | 3 | 3 | 1
| 1 | 2 | 3.024 | 3.27 | 3.5 | 3.8 | 3.2 | 3.2 | 3.3483 | 4 | 3 | 3.77 | 3.3398 | 4.31 | 3.71 | 3 | 3.55 | 5 | 3.25 | < 4 | 3.2 | 3 | 3 | 0
| 1 | 3 | 3.0975 | 3.36 | 3.3 | 3.4 | 3.5 | 3.3 | 3.4148 | 4 | 4 | 3.5 | 3.7831 | 3.39 | 3.76 | 3 | 3.65 | 4 | 3.324 | < 4 | 3.2 | 3 | 3 | 16
| 1 | 4 | 3.0835 | 3.54 | 3.4 | 3.5 | 3.4 | 3.5 | 3.708 | 5 | 3 | 3.78 | 3.3041 | 3.31 | 3.56 | 3 | 4.23 | 4 | 3.314 | < 4 | 3.1 | 3 | 3 | 3.3731
| 1 | 5 | 3.059 | 3.29 | 3.5 | 3.4 | 3.4 | 3.3 | 3.5935 | 5 | 3 | 3.52 | 3.5295 | 3.39 | 3.55 | 3 | 3.92 | 5 | 3.305 | < 4 | 3.1 | 3 | 3 | 15
| 1 | 6 | 3.2025 | 3.35 | 3.4 | 3.7 | 3.5 | 3.4 | 3.4859 | 5 | 3 | 3.67 | 3.2362 | 4.03 | 3.71 | 3 | 3.85 | 5 | 3.16 | < 4 | 3 | 3 | 3
| Mean |  | 3.0917 | 3.3617 | 3.4167 | 3.6833 | 3.4167 | 3.3333 | 3.498 | 4.6667 | 3.1667 | 3.6967 | 3.4485 | 3.73 | 3.7433 | 3 | 3.835 | 4.5 | 3.2827 |  | 3.1333 | 3 | 3
| Median |  | 3.0835 | 3.355 | 3.4 | 3.6 | 3.45 | 3.3 | 3.4617 | 5 | 3 | 3.72 | 3.419 | 3.67 | 3.71 | 3 | 3.83 | 4.5 | 3.3095 |  | 3.15 | 3 | 3
| Std Dev. |  | 0.060162 | 0.095376 | 0.075277 | 0.34303 | 0.1169 | 0.10328 | 0.13143 | 0.5164 | 0.40825 | 0.1686 | 0.19938 | 0.4201 | 0.22607 | 0 | 0.23645 | 0.54772 | 0.06774 |  | 0.08165 | 0 | 0
| Rel.Std.Dev. |  | 1.946% | 2.837% | 2.203% | 9.313% | 3.422% | 3.098% | 3.757% | 11.07% | 12.89% | 4.561% | 5.782% | 11.26% | 6.039% | 0% | 6.166% | 12.17% | 2.064% |  | 2.606% | 0% | 0%
| PDM3 |  | -8.344% | -0.3401% | 1.29% | 9.196% | 1.29% | -1.18% | 3.702% | 38.35% | -6.121% | 9.591% | 2.233% | 10.58% | 10.97% | -11.06% | 13.69% | 33.41% | -2.682% |  | -7.109% | -11.06% | -11.06%
| Z-Score (Absolute) |  | 0.8 | 0.22 | 0 | 0.66 | 0 | 0.21 | 0.2 | 3.08 | 0.62 | 0.69 | 0.08 | 0.77 | 0.6 | 1.03 | 1.03 | 2.67 | 0.33 | NA | 0.7 | 1.03 | 1.03
| Analytical results for Ce in OREAS 600 (Certified Value 47.7 ppm) | [22 empty] | Certified
Ce | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 13; 08: 15; 09: 17; 10: 19; 11: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*OES; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS | 1
2
| 1 | 1 | 01: 44.45; 02: 48.2848; 05: 47; 06: 48.29; 07: 55.1127; 08: 49.32; 09: 47; 10: 43.779; 11: 46.8 | 1
| 1 | 2 | 01: 43.61; 02: 48.0974; 05: 48; 06: 35.61; 07: 54.4101; 08: 48.25; 09: 49.4; 10: 44.361; 11: 46.2 | 0
| 1 | 3 | 01: 45.07; 02: 48.3682; 05: 50; 06: 48.11; 07: 56.2506; 08: 46.24; 09: 48.4; 10: 42.753; 11: 47.1 | 16
| 1 | 4 | 01: 45.1; 02: 47.3257; 05: 48; 06: 49.74; 07: 54.7589; 08: 48.74; 09: 49.1; 10: 43.994; 11: 46.4 | 47.7034
| 1 | 5 | 01: 44.68; 02: 48.278; 05: 49; 06: 48.08; 07: 54.6579; 08: 50.9; 09: 49.5; 10: 44.765; 11: 46.8 | 16
| 1 | 6 | 01: 44.31; 02: 47.5119; 05: 48; 06: 48.37; 07: 55.6382; 08: 49.2; 09: 50.1; 10: 43.863; 11: 47.9
| Mean |  | 01: 44.5367; 02: 47.9777; 05: 48.3333; 06: 46.3667; 07: 55.1381; 08: 48.775; 09: 48.9167; 10: 43.9192; 11: 46.8667
| Median |  | 01: 44.565; 02: 48.1877; 05: 48; 06: 48.2; 07: 54.9358; 08: 48.97; 09: 49.25; 10: 43.9285; 11: 46.8
| Std Dev. |  | 01: 0.55497; 02: 0.44572; 05: 1.0328; 06: 5.3061; 07: 0.69178; 08: 1.5295; 09: 1.0907; 10: 0.67809; 11: 0.59889
| Rel.Std.Dev. |  | 01: 1.246%; 02: 0.929%; 05: 2.137%; 06: 11.44%; 07: 1.255%; 08: 3.136%; 09: 2.23%; 10: 1.544%; 11: 1.278%
| PDM3 |  | 01: -6.638%; 02: 0.5749%; 05: 1.32%; 06: -2.802%; 07: 15.59%; 08: 2.246%; 09: 2.543%; 10: -7.933%; 11: -1.754%
| Z-Score (Absolute) |  | 01: 4.39; 02: 0.41; 05: 0; 06: 0.21; 07: 7.87; 08: 0.51; 09: 0.67; 10: 5.1; 11: 1.7
| Analytical results for Co in OREAS 600 (Certified Value 7.06 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 2
2
| 1 | 1 | 10.7486 | 7.3 | 7 | 8 | 6 | 6 | 6.7968 | 2 | 8 | 2.3 | 7.2862 | 6.8 | 7.55 | 6 | 8.1 | < 8 | 7.117 | 9 | 6 | 6 | 10 | 1
| 1 | 2 | 10.6123 | 7.3 | 7 | 8 | 7 | 6 | 6.8512 | 5 | 8 | 1.8 | 7.1859 | 7 | 7.79 | 6 | 7.1 | < 8 | 7.407 | 8 | 6 | 7 | 10 | 0
| 1 | 3 | 10.7646 | 7.5 | 7 | 8 | 7 | 7 | 6.7547 | 5 | 8 | 4.6 | 6.9488 | 6.7 | 7.63 | 5 | 7.8 | < 8 | 7.434 | 9 | 6 | 7 | 10 | 16
| 1 | 4 | 10.2809 | 7.4 | 7 | 8 | 7 | 6 | 6.7463 | 4 | 7 | 6.2 | 7.2847 | 6.6 | 7.96 | 5 | 8.6 | < 8 | 7.163 | 8 | 6 | 7 | 10 | 7.0554
| 1 | 5 | 10.5634 | 7.4 | 7 | 7 | 6 | 7 | 6.706 | < 1 | 8 | 6.2 | 7.0686 | 6.8 | 7.76 | 6 | 8.1 | < 8 | 7.301 | 9 | 7 | 7 | 10 | 17
| 1 | 6 | 10.8755 | 7.2 | 7 | 8 | 7 | 6 | 6.8033 | 2 | 8 | 6.3 | 7.3733 | 6.7 | 7.55 | 6 | 7.6 | < 8 | 7.231 | 9 | 6 | 8 | 10
| Mean |  | 10.6409 | 7.35 | 7 | 7.8333 | 6.6667 | 6.3333 | 6.7764 | 3.6 | 7.8333 | 4.5667 | 7.1912 | 6.7667 | 7.7067 | 5.6667 | 7.8833 |  | 7.2755 | 8.6667 | 6.1667 | 7 | 10
| Median |  | 10.6805 | 7.35 | 7 | 8 | 7 | 6 | 6.7757 | 4 | 8 | 5.4 | 7.2353 | 6.75 | 7.695 | 6 | 7.95 |  | 7.266 | 9 | 6 | 7 | 10
| Std Dev. |  | 0.20898 | 0.10488 | 0 | 0.40825 | 0.5164 | 0.5164 | 0.051123 | 1.5166 | 0.40825 | 2.0559 | 0.15793 | 0.13663 | 0.16058 | 0.5164 | 0.51153 |  | 0.1287 | 0.5164 | 0.40825 | 0.63246 | 0
| Rel.Std.Dev. |  | 1.964% | 1.427% | 0% | 5.212% | 7.746% | 8.154% | 0.7544% | 42.13% | 5.212% | 45.02% | 2.196% | 2.019% | 2.084% | 9.113% | 6.489% |  | 1.769% | 5.958% | 6.62% | 9.035% | 0%
| PDM3 |  | 50.82% | 4.176% | -0.7849% | 11.03% | -5.509% | -10.23% | -3.955% | -48.98% | 11.03% | -35.27% | 1.926% | -4.092% | 9.231% | -19.68% | 11.74% |  | 3.12% | 22.84% | -12.6% | -0.7849% | 41.74%
| Z-Score (Absolute) |  | 3.24 | 0.23 | 0.09 | 0.67 | 0.39 | 0.7 | 0.29 | 3.2 | 0.67 | 1.16 | 0.09 | 0.3 | 0.56 | 1.31 | 0.72 | NA | 0.16 | 1.44 | 0.85 | 0.09 | 2.65
| Analytical results for Cr in OREAS 600 (Certified Value 27.5 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES | 1
1
| 1 | 1 | 30.11 | 19 | 29 | 31 | 30 | 24 | 26.8477 | 32 | 28 | 35 | 31.7196 | 27 | 29.19 | 67 | 11 | 16 | 28.839 | 29 | 29 | 30 | < 50 | 1
| 1 | 2 | 32.1667 | 19 | 24 | 30 | 24 | 27 | 24.9493 | 33 | 29 | 20 | 33.695 | 28 | 30.88 | 62 | 12 | 16 | 24.786 | 31 | 20 | 30 | < 50 | 0
| 1 | 3 | 32.2725 | 24 | 26 | 27 | 29 | 26 | 30 | 30 | 27 | 24 | 29.0463 | 31 | 30.18 | 37 | 14 | 21 | 25.083 | 40 | 19 | 30 | < 50 | 16
| 1 | 4 | 32.61 | 23 | 24 | 26 | 46 | 25 | 25.195 | 30 | 30 | 20 | 33.0149 | 28 | 32 | 40 | 12 | 19 | 25.973 | 35 | 20 | 30 | 50 | 27.5108
| 1 | 5 | 30.6 | 23 | 24 | 29 | 28 | 28 | 29.2569 | 37 | 28 | 18 | 26.8357 | 26 | 31.19 | 42 | 11 | 16 | 24.919 | 41 | 24 | 30 | < 50 | 18
| 1 | 6 | 30.6533 | 24 | 24 | 32 | 31 | 25 | 28.9089 | 31 | 28 | 19 | 27.7049 | 26 | 30.9 | 44 | 15 | 17 | 22.372 | 33 | 20 | 40 | 50
| Mean |  | 31.4021 | 22 | 25.1667 | 29.1667 | 31.3333 | 25.8333 | 27.5263 | 32.1667 | 28.3333 | 22.6667 | 30.3361 | 27.6667 | 30.7233 | 48.6667 | 12.5 | 17.5 | 25.3287 | 34.8333 | 22 | 31.6667 | 50
| Median |  | 31.41 | 23 | 24 | 29.5 | 29.5 | 25.5 | 27.8783 | 31.5 | 28 | 20 | 30.383 | 27.5 | 30.89 | 43 | 12 | 16.5 | 25.001 | 34 | 20 | 30 | 50
| Std Dev. |  | 1.0653 | 2.3664 | 2.0412 | 2.3166 | 7.5807 | 1.472 | 2.1708 | 2.6394 | 1.0328 | 6.377 | 2.871 | 1.8619 | 0.95391 | 12.5804 | 1.6432 | 2.0736 | 2.0977 | 4.8339 | 3.8471 | 4.0825 | 0
| Rel.Std.Dev. |  | 3.393% | 10.76% | 8.111% | 7.943% | 24.19% | 5.698% | 7.886% | 8.206% | 3.645% | 28.13% | 9.464% | 6.73% | 3.105% | 25.85% | 13.15% | 11.85% | 8.282% | 13.88% | 17.49% | 12.89% | 0%
| PDM3 |  | 14.14% | -20.03% | -8.521% | 6.019% | 13.89% | -6.097% | 0.05645% | 16.92% | 2.99% | -17.61% | 10.27% | 0.5667% | 11.68% | 76.9% | -54.56% | -36.39% | -7.932% | 26.62% | -20.03% | 15.11% | 81.75%
| Z-Score (Absolute) |  | 0.67 | 1.39 | 0.86 | 0.18 | 0.01 | 0.55 | 0.18 | 0.84 | 0 | 1.79 | 0.44 | 0.15 | 0.53 | 4.47 | 3.48 | 2.53 | 0.66 | 1.16 | 1.7 | 0.37 | 4.76
| Analytical results for Cs in OREAS 600 (Certified Value 9.42 ppm) | [22 empty] | Certified
Cs | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 24; 12: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS | 2
3
| 1 | 1 | 01: 9.26; 02: 9.162; 05: 10.71; 06: 10.4339; 07: 9; 08: 9.46; 09: 9.38; 10: 8.524; 11: 9.3; 12: 8.5 | 1
| 1 | 2 | 01: 9.45; 02: 9.1567; 05: 9.08; 06: 10.0249; 07: 9.4; 08: 10.06; 09: 9.98; 10: 8.658; 11: 9.6; 12: 8.2 | 0
| 1 | 3 | 01: 9.62; 02: 9.2887; 05: 9.82; 06: 11.2342; 07: 9; 08: 9.75; 09: 9.66; 10: 8.458; 11: 9.5; 12: 8.4 | 16
| 1 | 4 | 01: 9.56; 02: 9.3111; 05: 10.53; 06: 10.1926; 07: 9.2; 08: 10.03; 09: 9.76; 10: 8.367; 11: 9.3; 12: 8.3 | 9.4182
| 1 | 5 | 01: 9.31; 02: 9.4569; 05: 9.98; 06: 10.3795; 07: 9.2; 08: 10.02; 09: 10; 10: 8.541; 11: 9.2; 12: 8.4 | 19
| 1 | 6 | 01: 9.34; 02: 9.1602; 05: 9.88; 06: 9.8888; 07: 9.2; 08: 9.81; 09: 9.9; 10: 8.602; 11: 9.7; 12: 8.5
| Mean |  | 01: 9.4233; 02: 9.2559; 05: 10; 06: 10.359; 07: 9.1667; 08: 9.855; 09: 9.78; 10: 8.525; 11: 9.4333; 12: 8.3833
| Median |  | 01: 9.395; 02: 9.2253; 05: 9.93; 06: 10.286; 07: 9.2; 08: 9.915; 09: 9.83; 10: 8.5325; 11: 9.4; 12: 8.4
| Std Dev. |  | 01: 0.14459; 02: 0.1203; 05: 0.57941; 06: 0.47588; 07: 0.15055; 08: 0.23193; 09: 0.23563; 10: 0.10327; 11: 0.19664; 12: 0.1169
| Rel.Std.Dev. |  | 01: 1.534%; 02: 1.3%; 05: 5.794%; 06: 4.594%; 07: 1.642%; 08: 2.353%; 09: 2.409%; 10: 1.211%; 11: 2.085%; 12: 1.394%
| PDM3 |  | 01: 0.05495%; 02: -1.722%; 05: 6.178%; 06: 9.989%; 07: -2.67%; 08: 4.638%; 09: 3.842%; 10: -9.483%; 11: 0.1611%; 12: -10.99%
| Z-Score (Absolute) |  | 01: 0.01; 02: 0.3; 05: 0.99; 06: 1.61; 07: 0.45; 08: 0.74; 09: 0.61; 10: 1.57; 11: 0.01; 12: 1.81
| Analytical results for Cu in OREAS 600 (Certified Value 482 ppm) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 505.3493 | 452.2 | 483 | 474 | 470 | 469 | 455.2517 | 524 | 497 | 483.5 | 487.1969 | 542 | 480.78 | 419 | 457 | 465 | 455.96 | 409 | 448 | 500 | 495 | 1
| 1 | 2 | 490.6162 | 451.7 | 495 | 505 | 460 | 448 | 452.0496 | 532 | 501 | 464.7 | 491.7695 | 532 | 510.24 | 415 | 463 | 478 | 446.74 | 469 | 461 | 516 | 490 | 0
| 1 | 3 | 492.3946 | 459.1 | 494 | 476 | 460 | 469 | 454.7566 | 562 | 500 | 473.1 | 497.8436 | 496 | 485.14 | 415 | 461 | 477 | 458 | 467 | 440 | 512 | 490 | 16
| 1 | 4 | 492.0048 | 461.2 | 511 | 483 | 460 | 477 | 462.7606 | 488 | 496 | 495.8 | 490.1896 | 482 | 497.47 | 420 | 469 | 476 | 450.66 | 478 | 456 | 508 | 495 | 481.9394
| 1 | 5 | 492.2983 | 446.1 | 477 | 483 | 474 | 469 | 459.3471 | 686 | 499 | 495.1 | 488.581 | 501 | 475 | 431 | 453 | 484 | 447.15 | 482 | 456 | 516 | 490 | 20
| 1 | 6 | 498.3754 | 453.4 | 505 | 488 | 458 | 492 | 452.7751 | 540 | 499 | 464.8 | 498.9718 | 528 | 492.66 | 447 | 459 | 490 | 499.84 | 484 | 455 | 528 | 495
| Mean |  | 495.1731 | 453.95 | 494.1667 | 484.8333 | 463.6667 | 470.6667 | 456.1568 | 555.3333 | 498.6667 | 479.5 | 492.4254 | 513.5 | 490.215 | 424.5 | 460.3333 | 478.3333 | 459.725 | 464.8333 | 452.6667 | 513.3333 | 492.5
| Median |  | 492.3465 | 452.8 | 494.5 | 483 | 460 | 469 | 455.0041 | 536 | 499 | 478.3 | 490.9796 | 514.5 | 488.9 | 419.5 | 460 | 477.5 | 453.31 | 473.5 | 455.5 | 514 | 492.5
| Std Dev. |  | 5.6674 | 5.4614 | 12.8128 | 11.125 | 6.6232 | 14.2642 | 4.12 | 68.43 | 1.8619 | 14.1502 | 4.894 | 23.7466 | 12.6955 | 12.486 | 5.465 | 8.4063 | 20.1756 | 28.1952 | 7.4744 | 9.3524 | 2.7386
| Rel.Std.Dev. |  | 1.145% | 1.203% | 2.593% | 2.295% | 1.428% | 3.031% | 0.9032% | 12.32% | 0.3734% | 2.951% | 0.9939% | 4.624% | 2.59% | 2.941% | 1.187% | 1.757% | 4.389% | 6.066% | 1.651% | 1.822% | 0.5561%
| PDM3 |  | 2.746% | -5.808% | 2.537% | 0.6005% | -3.792% | -2.339% | -5.35% | 15.23% | 3.471% | -0.5062% | 2.176% | 6.549% | 1.717% | -11.92% | -4.483% | -0.7483% | -4.609% | -3.549% | -6.074% | 6.514% | 2.191%
| Z-Score (Absolute) |  | 0.67 | 1.09 | 0.62 | 0.23 | 0.67 | 0.38 | 0.99 | 1.77 | 0.82 | 0 | 0.55 | 1.45 | 0.46 | 2.53 | 0.82 | 0.05 | 1.18 | 0.15 | 1.14 | 1.44 | 0.55
| Analytical results for Dy in OREAS 600 (Certified Value 2.05 ppm) | [22 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS | 2
3
| 1 | 1 | 01: 1.98; 02: 2.0767; 05: 2.1611; 06: 2.2; 07: 1.96; 08: 1.9 | 1
| 1 | 2 | 01: 1.94; 02: 2.0715; 05: 2.1714; 06: 2.22; 07: 2.07; 08: 1.9 | 0
| 1 | 3 | 01: 1.96; 02: 2.1233; 05: 2.1858; 06: 2.26; 07: 1.96; 08: 1.85 | 16
| 1 | 4 | 01: 2.04; 02: 2.0615; 05: 2.1172; 06: 2.2; 07: 2; 08: 1.85 | 2.0534
| 1 | 5 | 01: 1.99; 02: 2.1583; 05: 2.1274; 06: 2.05; 07: 2.04; 08: 1.85 | 21
| 1 | 6 | 01: 1.93; 02: 2.0796; 05: 2.1275; 06: 2.27; 07: 2.04; 08: 2
| Mean |  | 01: 1.9733; 02: 2.0951; 05: 2.1484; 06: 2.2; 07: 2.0117; 08: 1.8917
| Median |  | 01: 1.97; 02: 2.0782; 05: 2.1443; 06: 2.21; 07: 2.02; 08: 1.875
| Std Dev. |  | 01: 0.039833; 02: 0.037574; 05: 0.028064; 06: 0.079246; 07: 0.045789; 08: 0.058452
| Rel.Std.Dev. |  | 01: 2.019%; 02: 1.793%; 05: 1.306%; 06: 3.602%; 07: 2.276%; 08: 3.09%
| PDM3 |  | 01: -3.898%; 02: 2.035%; 05: 4.628%; 06: 7.141%; 07: -2.031%; 08: -7.875%
| Z-Score (Absolute) |  | 01: 0.62; 02: 0.32; 05: 0.73; 06: 1.13; 07: 0.32; 08: 1.25
| Analytical results for Er in OREAS 600 (Certified Value 0.97 ppm) | [22 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS | 2
3
| 1 | 1 | 01: 0.92; 02: 0.96788; 05: 0.43744; 06: 1.24; 07: 0.94; 08: 0.9 | 1
| 1 | 2 | 01: 0.92; 02: 0.97487; 05: 0.45114; 06: 1.07; 07: 0.98; 08: 0.85 | 0
| 1 | 3 | 01: 0.95; 02: 0.98394; 05: 0.41409; 06: 1.08; 07: 0.95; 08: 0.95 | 16
| 1 | 4 | 01: 0.95; 02: 0.96166; 05: 0.46113; 06: 1.08; 07: 0.97; 08: 0.85 | 0.96798
| 1 | 5 | 01: 0.91; 02: 0.92081; 05: 0.40575; 06: 1.08; 07: 0.99; 08: 0.9 | 22
| 1 | 6 | 01: 0.92; 02: 0.94835; 05: 0.45072; 06: 1.1; 07: 0.99; 08: 0.95
| Mean |  | 01: 0.92833; 02: 0.95958; 05: 0.43671; 06: 1.1083; 07: 0.97; 08: 0.9
| Median |  | 01: 0.92; 02: 0.96477; 05: 0.44408; 06: 1.08; 07: 0.975; 08: 0.9
| Std Dev. |  | 01: 0.017224; 02: 0.022491; 05: 0.022234; 06: 0.065243; 07: 0.020976; 08: 0.044721
| Rel.Std.Dev. |  | 01: 1.855%; 02: 2.344%; 05: 5.091%; 06: 5.887%; 07: 2.162%; 08: 4.969%
| PDM3 |  | 01: -4.096%; 02: -0.8677%; 05: -54.88%; 06: 14.5%; 07: 0.2083%; 08: -7.023%
| Z-Score (Absolute) |  | 01: 0.3; 02: 0.3; 05: 9.77; 06: 2.66; 07: 0.5; 08: 0.85
| Analytical results for Eu in OREAS 600 (Certified Value 1.03 ppm) | [22 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS | 2
2
| 1 | 1 | 01: 0.78; 02: 1.2983; 05: 1.2226; 06: 1.24; 07: 0.99; 08: 0.85 | 1
| 1 | 2 | 01: 0.75; 02: 1.2821; 05: 1.2202; 06: 1.11; 07: 1.05; 08: 0.85 | 0
| 1 | 3 | 01: 0.77; 02: 1.3294; 05: 1.2275; 06: 1.08; 07: 1.01; 08: 0.8 | 16
| 1 | 4 | 01: 0.77; 02: 1.2073; 05: 1.2427; 06: 1.1; 07: 1.04; 08: 0.8 | 1.0292
| 1 | 5 | 01: 0.78; 02: 1.0358; 05: 1.225; 06: 1.13; 07: 1.05; 08: 0.85 | 23
| 1 | 6 | 01: 0.77; 02: 1.06; 05: 1.2338; 06: 1.12; 07: 1.06; 08: 0.85
| Mean |  | 01: 0.77; 02: 1.2022; 05: 1.2286; 06: 1.13; 07: 1.0333; 08: 0.83333
| Median |  | 01: 0.77; 02: 1.2447; 05: 1.2263; 06: 1.115; 07: 1.045; 08: 0.85
| Std Dev. |  | 01: 0.010954; 02: 0.12631; 05: 0.0083265; 06: 0.056569; 07: 0.027325; 08: 0.02582
| Rel.Std.Dev. |  | 01: 1.423%; 02: 10.51%; 05: 0.6777%; 06: 5.006%; 07: 2.644%; 08: 3.098%
| PDM3 |  | 01: -25.19%; 02: 16.8%; 05: 19.37%; 06: 9.789%; 07: 0.3973%; 08: -19.03%
| Z-Score (Absolute) |  | 01: 1.4; 02: 0.61; 05: 0.74; 06: 0.17; 07: 0.17; 08: 1.11
| Analytical results for Fe in OREAS 600 (Certified Value 2.38 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES | 2
3
| 1 | 1 | 2.5116 | 2.35 | 2.32 | 2.31 | 2.34 | 2.31 | 2.326 | 2.4 | 2.41 | 2.08 | 2.4189 | 2.4 | 2.44 | 1.9 | 2.31 | 2.54 | 2.278 | 1.96 | 2.14 | 2.53 | 2.5 | 1
| 1 | 2 | 2.478 | 2.37 | 2.33 | 2.37 | 2.28 | 2.21 | 2.3574 | 2.42 | 2.43 | 1.95 | 2.4301 | 2.41 | 2.5 | 1.85 | 2.3 | 2.69 | 2.303 | 2.19 | 2.25 | 2.52 | 2.49 | 0
| 1 | 3 | 2.4066 | 2.36 | 2.33 | 2.25 | 2.26 | 2.34 | 2.4262 | 2.41 | 2.52 | 2.06 | 2.4441 | 2.39 | 2.48 | 1.85 | 2.38 | 2.63 | 2.248 | 2.19 | 2.22 | 2.51 | 2.54 | 16
| 1 | 4 | 2.4423 | 2.39 | 2.4 | 2.28 | 2.3 | 2.36 | 2.3357 | 2.46 | 2.48 | 2.24 | 2.4408 | 2.35 | 2.46 | 1.88 | 2.39 | 2.63 | 2.2 | 2.21 | 2.28 | 2.53 | 2.51 | 2.3757
| 1 | 5 | 2.3688 | 2.38 | 2.29 | 2.24 | 2.28 | 2.32 | 2.3651 | 2.45 | 2.48 | 2.22 | 2.4425 | 2.35 | 2.46 | 1.99 | 2.39 | 2.64 | 2.253 | 2.27 | 2.29 | 2.49 | 2.54 | 24
| 1 | 6 | 2.52 | 2.34 | 2.4 | 2.34 | 2.24 | 2.36 | 2.3074 | 2.48 | 2.46 | 2.13 | 2.4848 | 2.35 | 2.45 | 1.99 | 2.33 | 2.61 | 2.342 | 2.24 | 2.3 | 2.54 | 2.52
| Mean |  | 2.4545 | 2.365 | 2.345 | 2.2983 | 2.2833 | 2.3167 | 2.353 | 2.4367 | 2.4633 | 2.1133 | 2.4435 | 2.375 | 2.465 | 1.91 | 2.35 | 2.6233 | 2.2707 | 2.1767 | 2.2467 | 2.52 | 2.5167
| Median |  | 2.4602 | 2.365 | 2.33 | 2.295 | 2.28 | 2.33 | 2.3465 | 2.435 | 2.47 | 2.105 | 2.4416 | 2.37 | 2.46 | 1.89 | 2.355 | 2.63 | 2.2655 | 2.2 | 2.265 | 2.525 | 2.515
| Std Dev. |  | 0.059817 | 0.018708 | 0.045056 | 0.051153 | 0.034448 | 0.056095 | 0.041571 | 0.031411 | 0.039328 | 0.10801 | 0.022371 | 0.028107 | 0.021679 | 0.064807 | 0.041473 | 0.048854 | 0.049005 | 0.11057 | 0.059889 | 0.017889 | 0.020656
| Rel.Std.Dev. |  | 2.437% | 0.791% | 1.921% | 2.226% | 1.509% | 2.421% | 1.767% | 1.289% | 1.597% | 5.111% | 0.9155% | 1.183% | 0.8795% | 3.393% | 1.765% | 1.862% | 2.158% | 5.08% | 2.666% | 0.7099% | 0.8208%
| PDM3 |  | 3.319% | -0.4505% | -1.292% | -3.257% | -3.888% | -2.485% | -0.9572% | 2.566% | 3.689% | -11.04% | 2.855% | -0.02957% | 3.759% | -19.6% | -1.082% | 10.42% | -4.421% | -8.378% | -5.431% | 6.074% | 5.934%
| Z-Score (Absolute) |  | 0.76 | 0.09 | 0.06 | 0.41 | 0.52 | 0.11 | 0 | 0.62 | 0.82 | 1.54 | 0.67 | 0.16 | 0.83 | 3.3 | 0.02 | 2.01 | 0.61 | 0.99 | 0.79 | 1.24 | 1.22
| Analytical results for Ga in OREAS 600 (Certified Value 18.7 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 24 | 25 | [5 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 19.7125 | 17.56 | 20 | 20 | 20 | 20 | 17.4437 | 19 | 11.66 | 16.6066 | 18.2 | 18.4 | 17.5 | 18.575 | 17.4 | 15.8 | [5 empty] | 1
| 1 | 2 | 20.0475 | 17.43 | 20 | 20 | 20 | 20 | 17.6434 | 19 | 8.01 | 16.3114 | 18.3 | 19.4 | 18.5 | 19.143 | 18.2 | 16.8 | [5 empty] | 0
| 1 | 3 | 20.9833 | 17.86 | 20 | 20 | 20 | 20 | 18.5293 | 20 | 15.97 | 16.7314 | 17.5 | 19.33 | 17.7 | 18.983 | 18 | 16.2 | [5 empty] | 16
| 1 | 4 | 19.22 | 17.41 | 20 | 20 | 20 | 20 | 17.9445 | 19 | 19.35 | 16.7349 | 17.4 | 19.56 | 18.4 | 18.52 | 18 | 17.2 | [5 empty] | 18.7278
| 1 | 5 | 19.7975 | 17.47 | 20 | 20 | 20 | 20 | 18.2356 | 21 | 19.09 | 16.8359 | 18.2 | 19.71 | 18.6 | 18.775 | 18 | 17 | [5 empty] | 25
| 1 | 6 | 21.605 | 17.43 | 20 | 20 | 20 | 20 | 18.6328 | 21 | 18.88 | 16.5999 | 18.5 | 19.15 | 18.6 | 18.866 | 19 | 16.8
| Mean |  | 20.2276 | 17.5267 | 20 | 20 | 20 | 20 | 18.0716 | 19.8333 | 15.4933 | 16.6367 | 18.0167 | 19.2583 | 18.2167 | 18.8103 | 18.1 | 16.6333
| Median |  | 19.9225 | 17.45 | 20 | 20 | 20 | 20 | 18.0901 | 19.5 | 17.425 | 16.669 | 18.2 | 19.365 | 18.45 | 18.8205 | 18 | 16.8
| Std Dev. |  | 0.89066 | 0.17189 | 0 | 0 | 0 | 0 | 0.47855 | 0.98319 | 4.6947 | 0.1824 | 0.4535 | 0.46232 | 0.48751 | 0.23854 | 0.51769 | 0.52789
| Rel.Std.Dev. |  | 4.403% | 0.9807% | 0% | 0% | 0% | 0% | 2.648% | 4.957% | 30.3% | 1.096% | 2.517% | 2.401% | 2.676% | 1.268% | 2.86% | 3.174%
| PDM3 |  | 8.008% | -6.414% | 6.793% | 6.793% | 6.793% | 6.793% | -3.504% | 5.903% | -17.27% | -11.17% | -3.797% | 2.833% | -2.729% | 0.4405% | -3.352% | -11.18%
| Z-Score (Absolute) |  | 0.97 | 0.61 | 0.84 | 0.84 | 0.84 | 0.84 | 0.29 | 0.74 | 0.14 | 1.13 | 0.32 | 0.5 | 0.2 | 0.14 | 0.38 | 1.13
| Analytical results for Gd in OREAS 600 (Certified Value 3.12 ppm) | [22 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS | 2
2
| 1 | 1 | 01: 2.78; 02: 3.7087; 05: 3.5866; 06: 3.24; 07: 2.97; 08: 2.6 | 1
| 1 | 2 | 01: 2.73; 02: 3.5313; 05: 3.5964; 06: 3.2; 07: 3.09; 08: 2.4 | 0
| 1 | 3 | 01: 2.79; 02: 3.7126; 05: 3.7765; 06: 3.15; 07: 2.97; 08: 2.4 | 16
| 1 | 4 | 01: 2.81; 02: 3.6614; 05: 3.685; 06: 3.08; 07: 3; 08: 2.4 | 3.1244
| 1 | 5 | 01: 2.72; 02: 3.6853; 05: 3.6605; 06: 3.16; 07: 3.03; 08: 2.4 | 26
| 1 | 6 | 01: 2.72; 02: 3.6288; 05: 3.6749; 06: 3.26; 07: 3.07; 08: 2.6
| Mean |  | 01: 2.7583; 02: 3.6547; 05: 3.6633; 06: 3.1817; 07: 3.0217; 08: 2.4667
| Median |  | 01: 2.755; 02: 3.6734; 05: 3.6677; 06: 3.18; 07: 3.015; 08: 2.4
| Std Dev. |  | 01: 0.039707; 02: 0.068027; 05: 0.068957; 06: 0.065853; 07: 0.050761; 08: 0.10328
| Rel.Std.Dev. |  | 01: 1.44%; 02: 1.861%; 05: 1.882%; 06: 2.07%; 07: 1.68%; 08: 4.187%
| PDM3 |  | 01: -11.72%; 02: 16.97%; 05: 17.25%; 06: 1.833%; 07: -3.288%; 08: -21.05%
| Z-Score (Absolute) |  | 01: 0.52; 02: 0.83; 05: 0.85; 06: 0.12; 07: 0.12; 08: 0.96
| Analytical results for Ge in OREAS 600 (Indicative Value 0.78 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 14; 08: 15; 09: 17; 10: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: PF*MS | 2
2
| 1 | 1 | 01: 0.52; 02: 0.80553; 05: < 10; 06: 0.56; 07: 1; 08: 0.97; 09: 0.3; 10: < 20 | 1
| 1 | 2 | 01: 0.48; 02: 1.046; 05: < 10; 06: 0.53; 07: 0.9; 08: 1.24; 09: 0.4; 10: < 20 | 7
| 1 | 3 | 01: 0.39; 02: 0.94294; 05: < 10; 06: 0.4; 07: 1; 08: 1.31; 09: 0.8; 10: < 20 | 16
| 1 | 4 | 01: 0.42; 02: 0.91699; 05: < 10; 06: 0.37; 07: 1; 08: 1.2; 09: 0.8; 10: < 20 | 0.78035
| 1 | 5 | 01: 0.44; 02: 1.1286; 05: < 10; 06: 0.31; 07: 1; 08: 1.23; 09: 0.8; 10: < 20 | 9
| 1 | 6 | 01: 0.38; 02: 0.95254; 05: < 10; 06: 0.24; 07: 1; 08: 1.32; 09: 0.7; 10: < 20
| Mean |  | 01: 0.43833; 02: 0.96545; 06: 0.40167; 07: 0.98333; 08: 1.2117; 09: 0.63333
| Median |  | 01: 0.43; 02: 0.94774; 06: 0.385; 07: 1; 08: 1.235; 09: 0.75
| Std Dev. |  | 01: 0.053821; 02: 0.11111; 06: 0.12416; 07: 0.040825; 08: 0.12734; 09: 0.22509
| Rel.Std.Dev. |  | 01: 12.28%; 02: 11.51%; 06: 30.91%; 07: 4.152%; 08: 10.51%; 09: 35.54%
| PDM3 |  | 01: -43.83%; 02: 23.72%; 06: -48.53%; 07: 26.01%; 08: 55.27%; 09: -18.84%
| Z-Score (Absolute) |  | 01: 0.89; 02: 0.41; 05: NA; 06: 0.98; 07: 0.46; 08: 1.14; 09: 0.41; 10: NA
| Analytical results for Hf in OREAS 600 (Certified Value 2.9 ppm) | [22 empty] | Certified
Hf | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 24; 12: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS | 2
3
| 1 | 1 | 01: 3.1; 02: 2.9507; 05: 2.8; 06: 3.8575; 07: 3; 08: 2.94; 09: 2.44; 10: 2.88; 11: 3.6; 12: 2.4 | 1
| 1 | 2 | 01: 2.96; 02: 2.9286; 05: 1.9; 06: 3.6374; 07: 2.9; 08: 3.15; 09: 2.65; 10: 2.92; 11: 3; 12: 2.4 | 0
| 1 | 3 | 01: 2.85; 02: 2.9492; 05: 2.9; 06: 3.7951; 07: 3; 08: 2.96; 09: 2.61; 10: 2.957; 11: 2.8; 12: 2.4 | 16
| 1 | 4 | 01: 3.05; 02: 2.9417; 05: 3.1; 06: 3.5455; 07: 2.9; 08: 3; 09: 2.68; 10: 2.977; 11: 2.8; 12: 2.4 | 2.896
| 1 | 5 | 01: 2.93; 02: 2.9538; 05: 2.9; 06: 3.8175; 07: 3; 08: 3; 09: 2.71; 10: 2.927; 11: 2.6; 12: 2.4 | 27
| 1 | 6 | 01: 2.98; 02: 2.8113; 05: 2.9; 06: 3.7175; 07: 3; 08: 2.95; 09: 2.77; 10: 2.899; 11: 2.8; 12: 2.4
| Mean |  | 01: 2.9783; 02: 2.9225; 05: 2.75; 06: 3.7284; 07: 2.9667; 08: 3; 09: 2.6433; 10: 2.9267; 11: 2.9333; 12: 2.4
| Median |  | 01: 2.97; 02: 2.9454; 05: 2.9; 06: 3.7563; 07: 3; 08: 2.98; 09: 2.665; 10: 2.9235; 11: 2.8; 12: 2.4
| Std Dev. |  | 01: 0.088412; 02: 0.055236; 05: 0.42778; 06: 0.1191; 07: 0.05164; 08: 0.077717; 09: 0.11343; 10: 0.035892; 11: 0.35024; 12: 0
| Rel.Std.Dev. |  | 01: 2.969%; 02: 1.89%; 05: 15.56%; 06: 3.194%; 07: 1.741%; 08: 2.591%; 09: 4.291%; 10: 1.226%; 11: 11.94%; 12: 0%
| PDM3 |  | 01: 2.842%; 02: 0.9157%; 05: -5.042%; 06: 28.74%; 07: 2.439%; 08: 3.59%; 09: -8.726%; 10: 1.058%; 11: 1.288%; 12: -17.13%
| Z-Score (Absolute) |  | 01: 0.73; 02: 0.03; 05: 0.06; 06: 10.94; 07: 0.57; 08: 0.62; 09: 3.27; 10: 0.03; 11: 1.7; 12: 7.14
| Analytical results for Hg in OREAS 600 (Indicative Value < 1 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | 01: Lab | 1
| No | No | 01: 11 | ppm
|  |  | 01: 4A*OES | 2
2
| 1 | 1 | 01: < 1 | 1
| 1 | 2 | 01: < 1 | 4
| 1 | 3 | 01: < 1 | 16
| 1 | 4 | 01: < 1 | < 1
| 1 | 5 | 01: < 1 | 10
| 1 | 6 | 01: < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for Ho in OREAS 600 (Certified Value 0.36 ppm) | [22 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS | 2
3
| 1 | 1 | 01: 0.35; 02: 0.36258; 05: 0.39744; 06: 0.61; 07: 0.33; 08: 0.32 | 1
| 1 | 2 | 01: 0.35; 02: 0.37391; 05: 0.38114; 06: 0.44; 07: 0.36; 08: 0.32 | 0
| 1 | 3 | 01: 0.37; 02: 0.36376; 05: 0.41409; 06: 0.47; 07: 0.34; 08: 0.32 | 16
| 1 | 4 | 01: 0.36; 02: 0.3485; 05: 0.38113; 06: 0.46; 07: 0.34; 08: 0.32 | 0.35548
| 1 | 5 | 01: 0.35; 02: 0.35897; 05: 0.40575; 06: 0.43; 07: 0.35; 08: 0.32 | 28
| 1 | 6 | 01: 0.36; 02: 0.36644; 05: 0.38072; 06: 0.45; 07: 0.35; 08: 0.32
| Mean |  | 01: 0.35667; 02: 0.36236; 05: 0.39338; 06: 0.47667; 07: 0.345; 08: 0.32
| Median |  | 01: 0.355; 02: 0.36317; 05: 0.38929; 06: 0.455; 07: 0.345; 08: 0.32
| Std Dev. |  | 01: 0.008165; 02: 0.0084328; 05: 0.01455; 06: 0.066833; 07: 0.010488; 08: 0
| Rel.Std.Dev. |  | 01: 2.289%; 02: 2.327%; 05: 3.699%; 06: 14.02%; 07: 3.04%; 08: 0%
| PDM3 |  | 01: 0.3335%; 02: 1.935%; 05: 10.66%; 06: 34.09%; 07: -2.948%; 08: -9.981%
| Z-Score (Absolute) |  | 01: 0.08; 02: 0.08; 05: 0.94; 06: 2.52; 07: 0.4; 08: 1.1
| Analytical results for In in OREAS 600 (Certified Value 0.75 ppm) | [22 empty] | Certified
In | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 12; 08: 13; 09: 14; 10: 15; 11: 17; 12: 19; 13: 24; 14: 25 | ppm
|  |  | 01: 4A*OES; 02: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*OES/MS; 12: 4A*OES/MS; 13: 4A*MS; 14: 4A*MS | 2
3
| 1 | 1 | 01: < 5; 02: 0.76; 05: 0.73417; 06: < 10; 07: 0.85; 08: 0.69991; 09: 1.57; 10: 0.88; 11: 0.82; 12: 0.656; 13: 0.75; 14: 0.66 | 1
| 1 | 2 | 01: < 5; 02: 0.75; 05: 0.67691; 06: < 10; 07: 0.68; 08: 0.66959; 09: 1.52; 10: 0.8; 11: 0.77; 12: 0.661; 13: 0.8; 14: 0.68 | 0
| 1 | 3 | 01: < 5; 02: 0.71; 05: 0.72832; 06: < 10; 07: 0.79; 08: 0.74341; 09: 0.74; 10: 0.77; 11: 0.81; 12: 0.655; 13: 0.75; 14: 0.66 | 16
| 1 | 4 | 01: < 5; 02: 0.72; 05: 0.77845; 06: < 10; 07: 0.84; 08: 0.70753; 09: 0.76; 10: 0.85; 11: 0.89; 12: 0.656; 13: 0.75; 14: 0.68 | 0.75032
| 1 | 5 | 01: < 5; 02: 0.75; 05: 0.73767; 06: < 10; 07: 0.8; 08: 0.66478; 09: 0.81; 10: 0.8; 11: 0.87; 12: 0.655; 13: 0.75; 14: 0.62 | 29
| 1 | 6 | 01: < 5; 02: 0.75; 05: 0.81771; 06: < 10; 07: 0.81; 08: 0.70563; 09: 1.59; 10: 0.84; 11: 0.84; 12: 0.638; 13: 0.8; 14: 0.66
| Mean |  | 02: 0.74; 05: 0.74554; 07: 0.795; 08: 0.69847; 09: 1.165; 10: 0.82333; 11: 0.83333; 12: 0.6535; 13: 0.76667; 14: 0.66
| Median |  | 02: 0.75; 05: 0.73592; 07: 0.805; 08: 0.70277; 09: 1.165; 10: 0.82; 11: 0.83; 12: 0.6555; 13: 0.75; 14: 0.66
| Std Dev. |  | 02: 0.02; 05: 0.047934; 07: 0.06091; 08: 0.028721; 09: 0.4339; 10: 0.040332; 11: 0.043205; 12: 0.0079183; 13: 0.02582; 14: 0.021909
| Rel.Std.Dev. |  | 02: 2.703%; 05: 6.429%; 07: 7.662%; 08: 4.112%; 09: 37.24%; 10: 4.899%; 11: 5.185%; 12: 1.212%; 13: 3.368%; 14: 3.32%
| PDM3 |  | 02: -1.375%; 05: -0.6367%; 07: 5.955%; 08: -6.909%; 09: 55.27%; 10: 9.732%; 11: 11.06%; 12: -12.9%; 13: 2.179%; 14: -12.04%
| Z-Score (Absolute) |  | 01: NA; 02: 0.11; 05: 0.11; 06: NA; 07: 0.65; 08: 0.61; 09: 4.28; 10: 0.7; 11: 0.81; 12: 1.08; 13: 0.11; 14: 0.93
| Analytical results for K in OREAS 600 (Certified Value 1.8 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES |  | 2
3
| 1 | 1 | 1.85 | 1.8064 | 1.74 | 1.82 | 1.81 | 1.69 | 1.7003 | 1.86 | 2.05 | 1.8495 | 1.8362 | 1.81 | 1.75 | 1.79 | 1.98 | 1.87 | 1.86 | 1.66 | 1.87 | 1.9 |  | 1
| 1 | 2 | 1.775 | 1.8216 | 1.75 | 1.87 | 1.76 | 1.62 | 1.7347 | 1.82 | 1.66 | 1.8686 | 1.8363 | 1.82 | 1.72 | 1.78 | 2.18 | 1.856 | 1.54 | 1.75 | 1.85 | 1.8 |  | 0
| 1 | 3 | 1.735 | 1.7746 | 1.76 | 1.79 | 1.77 | 1.73 | 1.765 | 1.92 | 1.95 | 1.8528 | 1.8171 | 1.8 | 1.69 | 1.85 | 2.1 | 1.838 | 1.57 | 1.69 | 1.88 | 1.8 |  | 16
| 1 | 4 | 1.79 | 1.8281 | 1.79 | 1.8 | 1.78 | 1.73 | 1.7111 | 1.85 | 2.01 | 1.8751 | 1.7952 | 1.79 | 1.74 | 1.86 | 2.65 | 1.779 | 1.5 | 1.74 | 1.88 | 1.8 |  | 1.804
| 1 | 5 | 1.755 | 1.7969 | 1.71 | 1.78 | 1.74 | 1.7 | 1.7306 | 1.93 | 1.96 | 1.8714 | 1.817 | 1.81 | 1.8 | 1.85 | 2.75 | 1.84 | 1.48 | 1.76 | 1.85 | 1.8 |  | 30
| 1 | 6 | 1.815 | 1.7769 | 1.77 | 1.84 | 1.7 | 1.75 | 1.6867 | 1.89 | 1.91 | 1.9008 | 1.7924 | 1.79 | 1.75 | 1.81 | 2.06 | 1.991 | 1.46 | 1.74 | 1.85 | 1.8
| Mean |  | 1.7867 | 1.8008 | 1.7533 | 1.8167 | 1.76 | 1.7033 | 1.7214 | 1.8783 | 1.9233 | 1.8697 | 1.8157 | 1.8033 | 1.7417 | 1.8233 | 2.2867 | 1.8623 | 1.5683 | 1.7233 | 1.8633 | 1.8167
| Median |  | 1.7825 | 1.8016 | 1.755 | 1.81 | 1.765 | 1.715 | 1.7208 | 1.875 | 1.955 | 1.87 | 1.8171 | 1.805 | 1.745 | 1.83 | 2.14 | 1.848 | 1.52 | 1.74 | 1.86 | 1.8
| Std Dev. |  | 0.041553 | 0.022281 | 0.027325 | 0.033862 | 0.037417 | 0.046332 | 0.027992 | 0.042622 | 0.13794 | 0.018397 | 0.019034 | 0.012111 | 0.03656 | 0.034448 | 0.32813 | 0.07027 | 0.14838 | 0.039328 | 0.015055 | 0.040825
| Rel.Std.Dev. |  | 2.326% | 1.237% | 1.558% | 1.864% | 2.126% | 2.72% | 1.626% | 2.269% | 7.172% | 0.984% | 1.048% | 0.6716% | 2.099% | 1.889% | 14.35% | 3.773% | 9.461% | 2.282% | 0.808% | 2.247%
| PDM3 |  | -0.9597% | -0.1791% | -2.808% | 0.7032% | -2.438% | -5.579% | -4.578% | 4.122% | 6.616% | 3.643% | 0.6497% | -0.03586% | -3.454% | 1.073% | 26.76% | 3.235% | -13.06% | -4.471% | 3.29% | 0.7032%
| Z-Score (Absolute) |  | 0.28 | 0.11 | 0.69 | 0.09 | 0.61 | 1.3 | 1.08 | 0.84 | 1.81 | 0.74 | 0.08 | 0.08 | 0.83 | 0.17 | 5.85 | 0.33 | 3.67 | 1.06 | 0.66 | 0.09
| Analytical results for La in OREAS 600 (Certified Value 23 ppm) | [22 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 23 | 25 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 25.63 | 23.82 | 20 | 20 | 30 | 20 | 25.083 | 23 | 19.6 | 26.1636 | 25.3 | 19.9 | 25.6 | 21 | 22.3 | 21 | 25.7 |  |  |  |  | 1
| 1 | 2 | 24.925 | 23.85 | 20 | 20 | 20 | 20 | 25.1118 | 24 | 17.6 | 26.0474 | 26.55 | 19.4 | 27.1 | 22 | 22.53 | 19 | 25.3 |  |  |  |  | 0
| 1 | 3 | 24.615 | 24.14 | 20 | 20 | 20 | 20 | 25.124 | 25 | 20.7 | 26.4121 | 26.59 | 20.1 | 26.6 | 22 | 21.64 | 17 | 25.6 |  |  |  |  | 16
| 1 | 4 | 24.875 | 24.51 | 20 | 30 | 20 | 20 | 24.3101 | 23 | 20.9 | 26.4962 | 26.18 | 20 | 26.9 | 22 | 22.33 | 19 | 25.5 |  |  |  |  | 22.9505
| 1 | 5 | 24.635 | 24.03 | 20 | 20 | 20 | 20 | 25.4472 | 25 | 21.4 | 26.4895 | 26.39 | 21.5 | 27.3 | 22 | 22.55 | 17 | 25.7 |  |  |  |  | 31
| 1 | 6 | 25.55 | 23.82 | 20 | 20 | 20 | 20 | 24.5009 | 24 | 20.5 | 26.4197 | 26.19 | 22.4 | 27.4 | 22 | 22.02 | 17 | 26.2
| Mean |  | 25.0383 | 24.0283 | 20 | 21.6667 | 21.6667 | 20 | 24.9295 | 24 | 20.1167 | 26.3381 | 26.2 | 20.55 | 26.8167 | 21.8333 | 22.2283 | 18.3333 | 25.6667
| Median |  | 24.9 | 23.94 | 20 | 20 | 20 | 20 | 25.0974 | 24 | 20.6 | 26.4159 | 26.29 | 20.05 | 27 | 22 | 22.315 | 18 | 25.65
| Std Dev. |  | 0.44571 | 0.26947 | 0 | 4.0825 | 4.0825 | 0 | 0.43129 | 0.89443 | 1.3674 | 0.18711 | 0.47354 | 1.1467 | 0.66156 | 0.40825 | 0.34626 | 1.633 | 0.30111
| Rel.Std.Dev. |  | 1.78% | 1.121% | 0% | 18.84% | 18.84% | 0% | 1.73% | 3.727% | 6.797% | 0.7104% | 1.807% | 5.58% | 2.467% | 1.87% | 1.558% | 8.907% | 1.173%
| PDM3 |  | 9.097% | 4.696% | -12.86% | -5.594% | -5.594% | -12.86% | 8.623% | 4.573% | -12.35% | 14.76% | 14.16% | -10.46% | 16.85% | -4.868% | -3.147% | -20.12% | 11.83%
| Z-Score (Absolute) |  | 0.85 | 0.54 | 0.67 | 0.17 | 0.17 | 0.67 | 0.82 | 0.54 | 0.49 | 1.24 | 1.2 | 0.51 | 1.46 | 0.12 | 0 | 1.18 | 1.04
| Analytical results for Li in OREAS 600 (Certified Value 18.7 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | [6 empty] | 1
2
| 1 | 1 | 17.9133 | 19.4 | 20.4144 | 19 | 22.5 | 19.2881 | 17.6 | 19 | 18 | 18 | 31 | 16.783 | 38 | 20 | 18.5 | [6 empty] | 1
| 1 | 2 | 18.2725 | 18.6 | 20.4011 | 19 | 17.1 | 19.5475 | 17.7 | 20 | 17 | 17 | 33 | 16.514 | 42 | 20 | 18.5 | [6 empty] | 0
| 1 | 3 | 18.0425 | 19 | 20.2539 | 20 | 20.5 | 19.1856 | 17.8 | 21 | 17 | 17 | 32 | 16.259 | 43 | 20 | 18.5 | [6 empty] | 16
| 1 | 4 | 17.8833 | 19.1 | 19.8118 | 18 | 21.5 | 19.3438 | 17.8 | 19 | 17 | 17 | 33 | 15.688 | 45 | 20 | 18.5 | [6 empty] | 18.6871
| 1 | 5 | 18.45 | 18.6 | 19.7599 | 19 | 20.1 | 19.3916 | 17.8 | 20 | 18 | 18 | 34 | 16.274 | 46 | 20 | 18 | [6 empty] | 32
| 1 | 6 | 19.83 | 18.8 | 20.3068 | 19 | 20 | 19.2946 | 17.7 | 19 | 18 | 18 | 32 | 18.28 | 48 | 20 | 18.5
| Mean |  | 18.3986 | 18.9167 | 20.158 | 19 | 20.2833 | 19.3418 | 17.7333 | 19.6667 | 17.5 | 17.5 | 32.5 | 16.633 | 43.6667 | 20 | 18.4167
| Median |  | 18.1575 | 18.9 | 20.2803 | 19 | 20.3 | 19.3192 | 17.75 | 19.5 | 17.5 | 17.5 | 32.5 | 16.394 | 44 | 20 | 18.5
| Std Dev. |  | 0.7342 | 0.31252 | 0.29477 | 0.63246 | 1.8269 | 0.12189 | 0.08165 | 0.8165 | 0.54772 | 0.54772 | 1.0488 | 0.88443 | 3.5024 | 0 | 0.20412
| Rel.Std.Dev. |  | 3.991% | 1.652% | 1.462% | 3.329% | 9.007% | 0.6302% | 0.4604% | 4.152% | 3.13% | 3.13% | 3.227% | 5.317% | 8.021% | 0% | 1.108%
| PDM3 |  | -1.544% | 1.228% | 7.871% | 1.674% | 8.542% | 3.504% | -5.104% | 5.242% | -6.353% | -6.353% | 73.92% | -10.99% | 133.7% | 7.026% | -1.447%
| Z-Score (Absolute) |  | 0.52 | 0.05 | 0.67 | 0 | 0.75 | 0.2 | 0.74 | 0.39 | 0.87 | 0.87 | 7.86 | 1.57 | 14.36 | 0.58 | 0.34
| Analytical results for Lu in OREAS 600 (Certified Value 0.13 ppm) | [22 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab | [6 empty] | 1
| No | No | 01: 02; 02: 09; 09: 15; 11: 17; 12: 19; 13: 25 | [6 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 09: 4A*MS; 11: 4A*OES/MS; 12: 4A*OES/MS; 13: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 01: 0.12; 02: 0.12717; 09: 0.3; 11: 0.12; 12: 0.133; 13: 0.14 | [6 empty] | 1
| 1 | 2 | 01: 0.11; 02: 0.13684; 09: 0.2; 11: 0.12; 12: 0.14; 13: 0.12 | [6 empty] | 0
| 1 | 3 | 01: 0.12; 02: 0.13615; 09: 0.19; 11: 0.12; 12: 0.142; 13: 0.12 | [6 empty] | 16
| 1 | 4 | 01: 0.12; 02: 0.13578; 09: 0.25; 11: 0.12; 12: 0.139; 13: 0.12 | [6 empty] | 0.12628
| 1 | 5 | 01: 0.11; 02: 0.13694; 09: 0.2; 11: 0.12; 12: 0.135; 13: 0.12 | [6 empty] | 33
| 1 | 6 | 01: 0.12; 02: 0.12154; 09: 0.24; 11: 0.12; 12: 0.132; 13: 0.12
| Mean |  | 01: 0.11667; 02: 0.1324; 09: 0.23; 11: 0.12; 12: 0.13683; 13: 0.12333
| Median |  | 01: 0.12; 02: 0.13597; 09: 0.22; 11: 0.12; 12: 0.137; 13: 0.12
| Std Dev. |  | 01: 0.005164; 02: 0.0064986; 09: 0.041952; 11: 0; 12: 0.0040702; 13: 0.008165
| Rel.Std.Dev. |  | 01: 4.426%; 02: 4.908%; 09: 18.24%; 11: 0%; 12: 2.975%; 13: 6.62%
| PDM3 |  | 01: -7.614%; 02: 4.848%; 09: 82.13%; 11: -4.975%; 12: 8.355%; 13: -2.335%
| Z-Score (Absolute) |  | 01: 0.98; 02: 0.45; 09: 8.1; 11: 0.72; 12: 0.63; 13: 0.45
| Analytical results for Mg in OREAS 600 (Certified Value 0.772 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES | 3
3
| 1 | 1 | 0.8247 | 0.7564 | 0.74 | 0.76 | 0.74 | 0.71 | 0.73067 | 0.745 | 0.8 | 0.84 | 0.77689 | 0.7783 | 0.87 | 0.71 | 0.77 | 0.74 | 0.782 | 0.75 | 0.74 | 0.82 | 0.81 | 1
| 1 | 2 | 0.8225 | 0.7674 | 0.74 | 0.78 | 0.72 | 0.68 | 0.72875 | 0.74 | 0.81 | 0.81 | 0.78001 | 0.776 | 0.89 | 0.7 | 0.75 | 0.78 | 0.761 | 0.82 | 0.76 | 0.82 | 0.81 | 0
| 1 | 3 | 0.7976 | 0.7664 | 0.74 | 0.76 | 0.71 | 0.73 | 0.74823 | 0.73 | 0.83 | 0.81 | 0.78457 | 0.7636 | 0.89 | 0.71 | 0.79 | 0.76 | 0.78 | 0.83 | 0.73 | 0.82 | 0.82 | 16
| 1 | 4 | 0.8084 | 0.7706 | 0.76 | 0.75 | 0.73 | 0.73 | 0.7519 | 0.765 | 0.81 | 0.84 | 0.78109 | 0.7662 | 0.89 | 0.71 | 0.8 | 0.76 | 0.752 | 0.77 | 0.75 | 0.82 | 0.82 | 0.77164
| 1 | 5 | 0.7906 | 0.7636 | 0.73 | 0.75 | 0.74 | 0.72 | 0.7488 | 0.74 | 0.82 | 0.82 | 0.77995 | 0.7723 | 0.87 | 0.73 | 0.79 | 0.78 | 0.766 | 0.82 | 0.75 | 0.79 | 0.82 | 34
| 1 | 6 | 0.82667 | 0.7595 | 0.76 | 0.77 | 0.73 | 0.73 | 0.74361 | 0.775 | 0.81 | 0.79 | 0.78345 | 0.7655 | 0.87 | 0.73 | 0.77 | 0.76 | 0.821 | 0.82 | 0.76 | 0.82 | 0.81
| Mean |  | 0.81174 | 0.76398 | 0.745 | 0.76167 | 0.72833 | 0.71667 | 0.742 | 0.74917 | 0.81333 | 0.81833 | 0.78099 | 0.77032 | 0.88 | 0.715 | 0.77833 | 0.76333 | 0.777 | 0.80167 | 0.74833 | 0.815 | 0.815
| Median |  | 0.81545 | 0.765 | 0.74 | 0.76 | 0.73 | 0.725 | 0.74592 | 0.7425 | 0.81 | 0.815 | 0.78055 | 0.76925 | 0.88 | 0.71 | 0.78 | 0.76 | 0.773 | 0.82 | 0.75 | 0.82 | 0.815
| Std Dev. |  | 0.015261 | 0.0052735 | 0.012247 | 0.01169 | 0.01169 | 0.019664 | 0.0098944 | 0.017151 | 0.010328 | 0.019408 | 0.0027468 | 0.0060865 | 0.010954 | 0.012247 | 0.018348 | 0.015055 | 0.02438 | 0.033116 | 0.01169 | 0.012247 | 0.0054772
| Rel.Std.Dev. |  | 1.88% | 0.6903% | 1.644% | 1.535% | 1.605% | 2.744% | 1.333% | 2.289% | 1.27% | 2.372% | 0.3517% | 0.7901% | 1.245% | 1.713% | 2.357% | 1.972% | 3.138% | 4.131% | 1.562% | 1.503% | 0.6721%
| PDM3 |  | 5.197% | -0.9927% | -3.453% | -1.293% | -5.613% | -7.125% | -3.842% | -2.913% | 5.403% | 6.051% | 1.212% | -0.1719% | 14.04% | -7.341% | 0.867% | -1.077% | 0.6942% | 3.891% | -3.021% | 5.619% | 5.619%
| Z-Score (Absolute) |  | 0.99 | 0.15 | 0.6 | 0.21 | 1 | 1.1 | 0.67 | 0.5 | 1.02 | 1.14 | 0.25 | 0 | 2.61 | 1.32 | 0.19 | 0.17 | 0.16 | 0.99 | 0.52 | 1.06 | 1.06
| Analytical results for Mn in OREAS 600 (Certified Value 0.071 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.077612 | 0.0675 | 0.0713 | 0.0764 | 0.071 | 0.0667 | 0.070322 | 0.0711 | 0.0731 | 0.0661 | 0.074341 | 0.0761 | 0.0741 | 0.0636 | 0.0637 | 0.07 | 0.071823 | 0.0612 | 0.0728 | 0.0744 | 0.0718 | 1
| 1 | 2 | 0.075103 | 0.0682 | 0.0714 | 0.0779 | 0.0704 | 0.0662 | 0.070783 | 0.0711 | 0.0751 | 0.0604 | 0.075432 | 0.0754 | 0.0734 | 0.0614 | 0.064 | 0.07 | 0.072104 | 0.0684 | 0.0701 | 0.0752 | 0.0716 | 0
| 1 | 3 | 0.076414 | 0.068 | 0.072 | 0.0748 | 0.07 | 0.0702 | 0.071287 | 0.0709 | 0.0758 | 0.0682 | 0.073808 | 0.0754 | 0.0736 | 0.0626 | 0.0645 | 0.07 | 0.070425 | 0.0681 | 0.0686 | 0.076 | 0.0704 | 16
| 1 | 4 | 0.07676 | 0.0693 | 0.0734 | 0.0744 | 0.0697 | 0.069 | 0.071499 | 0.0718 | 0.0753 | 0.0703 | 0.075226 | 0.0734 | 0.0736 | 0.0628 | 0.0641 | 0.07 | 0.068618 | 0.0691 | 0.0693 | 0.075 | 0.0698 | 0.071033
| 1 | 5 | 0.076011 | 0.0673 | 0.0701 | 0.0745 | 0.0703 | 0.0697 | 0.071223 | 0.0703 | 0.0748 | 0.0697 | 0.075315 | 0.0746 | 0.0744 | 0.0651 | 0.0638 | 0.07 | 0.070554 | 0.0705 | 0.0704 | 0.0736 | 0.0704 | 35
| 1 | 6 | 0.077412 | 0.067 | 0.0725 | 0.0771 | 0.0703 | 0.069 | 0.070289 | 0.0728 | 0.0753 | 0.0663 | 0.076053 | 0.0738 | 0.0734 | 0.0674 | 0.0648 | 0.07 | 0.071916 | 0.0696 | 0.0701 | 0.0744 | 0.0712
| Mean |  | 0.076552 | 0.067883 | 0.071783 | 0.07585 | 0.070283 | 0.068467 | 0.070901 | 0.071333 | 0.0749 | 0.066833 | 0.075029 | 0.074783 | 0.07375 | 0.063817 | 0.06415 | 0.07 | 0.070907 | 0.067817 | 0.070217 | 0.074767 | 0.070867
| Median |  | 0.076587 | 0.06775 | 0.0717 | 0.0756 | 0.0703 | 0.069 | 0.071003 | 0.0711 | 0.0752 | 0.06725 | 0.075271 | 0.075 | 0.0736 | 0.0632 | 0.06405 | 0.07 | 0.071188 | 0.06875 | 0.0701 | 0.0747 | 0.0708
| Std Dev. |  | 0.00092932 | 0.00082321 | 0.0011303 | 0.0014896 | 0.00043551 | 0.0016342 | 0.0005167 | 0.0008641 | 0.00094021 | 0.0035865 | 0.00081183 | 0.00104 | 0.00040866 | 0.0021395 | 0.00042308 | 0 | 0.0013331 | 0.0033535 | 0.0014275 | 0.00082381 | 0.00078655
| Rel.Std.Dev. |  | 1.214% | 1.213% | 1.575% | 1.964% | 0.6196% | 2.387% | 0.7288% | 1.211% | 1.255% | 5.366% | 1.082% | 1.391% | 0.5541% | 3.353% | 0.6595% | 0% | 1.88% | 4.945% | 2.033% | 1.102% | 1.11%
| PDM3 |  | 7.77% | -4.434% | 1.056% | 6.781% | -1.056% | -3.613% | -0.1866% | 0.4226% | 5.444% | -5.912% | 5.626% | 5.28% | 3.825% | -10.16% | -9.69% | -1.454% | -0.178% | -4.528% | -1.149% | 5.256% | -0.2343%
| Z-Score (Absolute) |  | 1.34 | 0.71 | 0.21 | 1.17 | 0.15 | 0.58 | 0 | 0.1 | 0.95 | 0.96 | 0.98 | 0.92 | 0.67 | 1.68 | 1.6 | 0.21 | 0 | 0.42 | 0.28 | 0.91 | 0.01
| Analytical results for Mo in OREAS 600 (Certified Value 2.2 ppm) | [22 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | 2.31 | 2.2 | 2 | 1 | 1 | 1 | 2.0027 | 3 | 3 | 2.15 | 2.6 | 2.44 | < 2.5 | 4.45 | < 3 | 2.239 | < 2 | 2.5 | 2.5 |  |  | 1
| 1 | 2 | 2.0781 | 2.1 | 1 | 1 | 1 | 1 | 1.9608 | 3 | 3 | 1.69 | 2.5 | 2.02 | 4 | 2.93 | < 3 | 2.189 | < 2 | 2.5 | 2.5 |  |  | 0
| 1 | 3 | 2.012 | 2.1 | 1 | 1 | 1 | 2 | 1.9705 | 2 | 3 | 1.91 | 2.1 | 2.01 | 3 | 2.85 | < 3 | 2.203 | < 2 | 2.5 | 2.5 |  |  | 16
| 1 | 4 | 2.1884 | 2 | 2 | 2 | 5 | 2 | 1.8537 | 3 | 2 | 2.01 | 2.2 | 2.37 | < 2.5 | 2.86 | < 3 | 2.301 | < 2 | 2.5 | 2.5 |  |  | 2.2033
| 1 | 5 | 2.1318 | 2 | 2 | 1 | 1 | 1 | 2.1007 | 2 | 3 | 1.85 | 2.1 | 2.18 | < 2.5 | 2.77 | < 3 | 2.18 | < 2 | 2.5 | 2.5 |  |  | 36
| 1 | 6 | 2.1801 | 2 | 2 | 2 | 1 | 2 | 1.9309 | 3 | 2 | 1.72 | 2.2 | 2.02 | 4 | 2.71 | < 3 | 2.053 | 2 | 3 | 1.5
| Mean |  | 2.1501 | 2.0667 | 1.6667 | 1.3333 | 1.6667 | 1.5 | 1.9699 | 2.6667 | 2.6667 | 1.8883 | 2.2833 | 2.1733 | 3.6667 | 3.095 |  | 2.1942 | 2 | 2.5833 | 2.3333
| Median |  | 2.156 | 2.05 | 2 | 1 | 1 | 1.5 | 1.9656 | 3 | 3 | 1.88 | 2.2 | 2.1 | 4 | 2.855 |  | 2.196 | 2 | 2.5 | 2.5
| Std Dev. |  | 0.10249 | 0.08165 | 0.5164 | 0.5164 | 1.633 | 0.54772 | 0.081597 | 0.5164 | 0.5164 | 0.17486 | 0.2137 | 0.19159 | 0.57735 | 0.66818 |  | 0.082045 |  | 0.20412 | 0.40825
| Rel.Std.Dev. |  | 4.767% | 3.951% | 30.98% | 38.73% | 97.98% | 36.51% | 4.142% | 19.36% | 19.36% | 9.26% | 9.359% | 8.815% | 15.75% | 21.59% |  | 3.739% |  | 7.902% | 17.5%
| PDM3 |  | -2.416% | -6.202% | -24.36% | -39.49% | -24.36% | -31.92% | -10.6% | 21.03% | 21.03% | -14.3% | 3.632% | -1.361% | 66.42% | 40.47% |  | -0.4154% | -9.228% | 17.25% | 5.901%
| Z-Score (Absolute) |  | 0.02 | 0.19 | 0.99 | 1.65 | 2.32 | 1.32 | 0.38 | 1.01 | 1.01 | 0.54 | 0.24 | 0.02 | 3 | 1.32 | NA | 0.06 | 0.32 | 0.67 | 0.67
| Analytical results for Na in OREAS 600 (Certified Value 0.587 wt.%) | [22 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.65953 | 0.5793 | 0.56 | 0.6 | 0.58 | 0.56 | 0.58398 | 0.59 | 0.64 | 0.58831 | 0.5808 | 0.61 | 0.57 | 0.57 | 0.62 | 0.601 | 0.53 | 0.61 | 0.62 |  |  | 1
| 1 | 2 | 0.65867 | 0.5884 | 0.56 | 0.62 | 0.55 | 0.52 | 0.58803 | 0.6 | 0.59 | 0.58849 | 0.5828 | 0.6 | 0.56 | 0.56 | 0.65 | 0.589 | 0.54 | 0.61 | 0.61 |  |  | 0
| 1 | 3 | 0.65833 | 0.58 | 0.57 | 0.6 | 0.56 | 0.56 | 0.5926 | 0.61 | 0.62 | 0.58125 | 0.5778 | 0.59 | 0.54 | 0.58 | 0.64 | 0.581 | 0.52 | 0.61 | 0.61 |  |  | 16
| 1 | 4 | 0.6712 | 0.5875 | 0.58 | 0.59 | 0.55 | 0.56 | 0.5976 | 0.59 | 0.63 | 0.58819 | 0.5711 | 0.59 | 0.55 | 0.59 | 0.64 | 0.571 | 0.52 | 0.61 | 0.62 |  |  | 0.58651
| 1 | 5 | 0.6588 | 0.5892 | 0.56 | 0.59 | 0.57 | 0.55 | 0.59899 | 0.6 | 0.62 | 0.58713 | 0.5758 | 0.6 | 0.58 | 0.59 | 0.65 | 0.591 | 0.54 | 0.59 | 0.62 |  |  | 37
| 1 | 6 | 0.677 | 0.561 | 0.58 | 0.6 | 0.56 | 0.56 | 0.59603 | 0.59 | 0.6 | 0.59018 | 0.5692 | 0.6 | 0.56 | 0.58 | 0.63 | 0.661 | 0.53 | 0.6 | 0.61
| Mean |  | 0.66392 | 0.5809 | 0.56833 | 0.6 | 0.56167 | 0.55167 | 0.59287 | 0.59667 | 0.61667 | 0.58726 | 0.57625 | 0.59833 | 0.56 | 0.57833 | 0.63833 | 0.599 | 0.53 | 0.605 | 0.615
| Median |  | 0.65917 | 0.58375 | 0.565 | 0.6 | 0.56 | 0.56 | 0.59432 | 0.595 | 0.62 | 0.58825 | 0.5768 | 0.6 | 0.56 | 0.58 | 0.64 | 0.59 | 0.53 | 0.61 | 0.615
| Std Dev. |  | 0.0081037 | 0.010659 | 0.0098319 | 0.010954 | 0.01169 | 0.016021 | 0.00587 | 0.008165 | 0.018619 | 0.0031027 | 0.0053373 | 0.0075277 | 0.014142 | 0.01169 | 0.01169 | 0.032 | 0.0089443 | 0.0083666 | 0.0054772
| Rel.Std.Dev. |  | 1.221% | 1.835% | 1.73% | 1.826% | 2.081% | 2.904% | 0.9901% | 1.368% | 3.019% | 0.5283% | 0.9262% | 1.258% | 2.525% | 2.021% | 1.831% | 5.342% | 1.688% | 1.383% | 0.8906%
| PDM3 |  | 13.2% | -0.9566% | -3.099% | 2.3% | -4.236% | -5.941% | 1.085% | 1.732% | 5.142% | 0.1271% | -1.749% | 2.016% | -4.52% | -1.394% | 8.836% | 2.129% | -9.635% | 3.152% | 4.857%
| Z-Score (Absolute) |  | 2.73 | 0.08 | 0.67 | 0.45 | 0.91 | 1.04 | 0.2 | 0.34 | 1.05 | 0 | 0.39 | 0.39 | 0.97 | 0.32 | 1.82 | 0.02 | 2.04 | 0.74 | 0.99
| Analytical results for Nb in OREAS 600 (Certified Value 7.98 ppm) | [22 empty] | Certified
Nb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 12; 08: 14; 09: 15; 11: 17; 12: 19; 13: 23; 14: 25 |  |  | ppm
|  |  | 01: 4A*OES; 02: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES; 11: 4A*OES/MS; 12: 4A*OES/MS; 13: 4A*OES; 14: 4A*MS |  |  | 2
3
| 1 | 1 | 01: 8.25; 02: 9.85; 05: 7.7581; 06: < 10; 07: 6.6; 08: 7.8; 09: 8; 11: 7.7; 12: 9.005; 13: 8; 14: 8 |  |  | 1
| 1 | 2 | 01: 8.09; 02: 8.82; 05: 7.6535; 06: < 10; 07: 5; 08: 7.7; 09: 8; 11: 8.6; 12: 8.732; 13: 6; 14: 8.5 |  |  | 8
| 1 | 3 | 01: 7.905; 02: 8.33; 05: 7.6954; 06: < 10; 07: 7.7; 08: 7.6; 09: 8; 11: 7.9; 12: 8.259; 13: 7; 14: 8.5 |  |  | 16
| 1 | 4 | 01: 7.605; 02: 8.15; 05: 7.7143; 06: < 10; 07: 8.5; 08: 7.5; 09: 8; 11: 7.9; 12: 8.698; 13: 7; 14: 8.5 |  |  | 7.9789
| 1 | 5 | 01: 7.2; 02: 8.27; 05: 7.6148; 06: < 10; 07: 8.4; 08: 7.6; 09: 8; 11: 8.1; 12: 8.653; 13: 7; 14: 8.5 |  |  | 38
| 1 | 6 | 01: 8.175; 02: 8.23; 05: 7.6649; 06: < 10; 07: 8.1; 08: 7.7; 09: 8; 11: 8.2; 12: 8.741; 13: 7; 14: 8.5
| Mean |  | 01: 7.8708; 02: 8.6083; 05: 7.6835; 07: 7.3833; 08: 7.65; 09: 8; 11: 8.0667; 12: 8.6813; 13: 7; 14: 8.4167
| Median |  | 01: 7.9975; 02: 8.3; 05: 7.6802; 07: 7.9; 08: 7.65; 09: 8; 11: 8; 12: 8.715; 13: 7; 14: 8.5
| Std Dev. |  | 01: 0.40154; 02: 0.65294; 05: 0.050255; 07: 1.3556; 08: 0.10488; 09: 0; 11: 0.31411; 12: 0.24097; 13: 0.63246; 14: 0.20412
| Rel.Std.Dev. |  | 01: 5.102%; 02: 7.585%; 05: 0.6541%; 07: 18.36%; 08: 1.371%; 09: 0%; 11: 3.894%; 12: 2.776%; 13: 9.035%; 14: 2.425%
| PDM3 |  | 01: -1.354%; 02: 7.889%; 05: -3.702%; 07: -7.464%; 08: -4.122%; 09: 0.2644%; 11: 1.1%; 12: 8.804%; 13: -12.27%; 14: 5.487%
| Z-Score (Absolute) |  | 01: 0.16; 02: 1.07; 05: 0.63; 06: NA; 07: 0.19; 08: 0.72; 09: 0.16; 11: 0.33; 12: 1.87; 13: 1.85; 14: 1.21
| Analytical results for Nd in OREAS 600 (Certified Value 19.5 ppm) | [22 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  |  | 1
2
| 1 | 1 | 01: 18.2; 02: 20.1147; 05: 20.896; 06: 18.79; 07: 19.2; 08: 19.5 |  |  | 1
| 1 | 2 | 01: 17.82; 02: 20.1939; 05: 20.4892; 06: 20.37; 07: 17.6; 08: 19 |  |  | 0
| 1 | 3 | 01: 18.01; 02: 20.0185; 05: 20.1424; 06: 19.96; 07: 18.2; 08: 19.5 |  |  | 16
| 1 | 4 | 01: 18.52; 02: 19.3765; 05: 20.6763; 06: 20.26; 07: 18.9; 08: 18.7 |  |  | 19.4882
| 1 | 5 | 01: 18.18; 02: 19.5745; 05: 20.8529; 06: 20.41; 07: 19.2; 08: 19.6 |  |  | 39
| 1 | 6 | 01: 18.13; 02: 19.9662; 05: 21.168; 06: 20.29; 07: 19; 08: 19.3
| Mean |  | 01: 18.1433; 02: 19.874; 05: 20.7041; 06: 20.0133; 07: 18.6833; 08: 19.2667
| Median |  | 01: 18.155; 02: 19.9923; 05: 20.7646; 06: 20.275; 07: 18.95; 08: 19.4
| Std Dev. |  | 01: 0.23192; 02: 0.3246; 05: 0.35683; 06: 0.61989; 07: 0.64627; 08: 0.35024
| Rel.Std.Dev. |  | 01: 1.278%; 02: 1.633%; 05: 1.723%; 06: 3.097%; 07: 3.459%; 08: 1.818%
| PDM3 |  | 01: -6.901%; 02: 1.98%; 05: 6.239%; 06: 2.694%; 07: -4.13%; 08: -1.137%
| Z-Score (Absolute) |  | 01: 1.22; 02: 0.26; 05: 0.97; 06: 0.59; 07: 0.76; 08: 0.26
| Analytical results for Ni in OREAS 600 (Certified Value 16.5 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 1
1
| 1 | 1 | 20.2436 | 14.3 | 17 | 19 | 13 | 16 | 14.8083 | 48 | 17 | 13.9 | 19.6535 | 14 | 15.6 | 44 | 20.5 | 12 | 13.53 | 16 | 17 | 20 | 20 | 1
| 1 | 2 | 20.0144 | 13.1 | 16 | 15 | 12 | 16 | 14.0144 | 61 | 15 | 7.7 | 21.7111 | 15 | 17.68 | 31 | 16.5 | 11 | 14.5 | 16 | 16 | 20 | 15 | 0
| 1 | 3 | 17.7984 | 14.3 | 17 | 16 | 13 | 16 | 16.4927 | 48 | 16 | 15.5 | 20.7999 | 16 | 16.64 | 23 | 20.8 | 12 | 15.337 | 21 | 15 | 20 | 20 | 16
| 1 | 4 | 20.0508 | 15.1 | 15 | 15 | 28 | 16 | 13.3611 | 44 | 18 | 14.9 | 21.9148 | 15 | 15.6 | 19 | 22.5 | 12 | 13.01 | 19 | 15 | 15 | 20 | 16.5011
| 1 | 5 | 18.2063 | 14.8 | 16 | 14 | 17 | 17 | 15.183 | 54 | 16 | 14.7 | 21.2042 | 14 | 17.68 | 20 | 21.2 | 11 | 14.63 | 22 | 17 | 20 | 20 | 40
| 1 | 6 | 19.5768 | 14.9 | 16 | 16 | 17 | 16 | 13.7112 | 53 | 15 | 16.1 | 21.1497 | 14 | 16.64 | 26 | 26.1 | 14 | 13.57 | 18 | 14 | 25 | 15
| Mean |  | 19.315 | 14.4167 | 16.1667 | 15.8333 | 16.6667 | 16.1667 | 14.5951 | 51.3333 | 16.1667 | 13.8 | 21.0722 | 14.6667 | 16.64 | 27.1667 | 21.2667 | 12 | 14.0962 | 18.6667 | 15.6667 | 20 | 18.3333
| Median |  | 19.7956 | 14.55 | 16 | 15.5 | 15 | 16 | 14.4113 | 50.5 | 16 | 14.8 | 21.177 | 14.5 | 16.64 | 24.5 | 21 | 12 | 14.035 | 18.5 | 15.5 | 20 | 20
| Std Dev. |  | 1.0479 | 0.72226 | 0.75277 | 1.7224 | 5.9554 | 0.40825 | 1.1518 | 5.9889 | 1.169 | 3.0796 | 0.80355 | 0.8165 | 0.9302 | 9.3256 | 3.1117 | 1.0954 | 0.86774 | 2.5033 | 1.2111 | 3.1623 | 2.582
| Rel.Std.Dev. |  | 5.425% | 5.01% | 4.656% | 10.88% | 35.73% | 2.525% | 7.892% | 11.67% | 7.231% | 22.32% | 3.813% | 5.567% | 5.59% | 34.33% | 14.63% | 9.129% | 6.156% | 13.41% | 7.73% | 15.81% | 14.08%
| PDM3 |  | 17.05% | -12.63% | -2.027% | -4.047% | 1.003% | -2.027% | -11.55% | 211.1% | -2.027% | -16.37% | 27.7% | -11.12% | 0.8415% | 64.64% | 28.88% | -27.28% | -14.57% | 13.12% | -5.057% | 21.2% | 11.1%
| Z-Score (Absolute) |  | 1.2 | 0.67 | 0 | 0.13 | 0.67 | 0 | 0.6 | 13.42 | 0 | 0.44 | 1.87 | 0.57 | 0.18 | 4.2 | 1.58 | 1.59 | 0.79 | 0.95 | 0.19 | 1.46 | 0.83
| Analytical results for P in OREAS 600 (Certified Value 0.06 wt.%) | [22 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.060619 | 0.0595 | 0.058 | 0.059 | 0.061 | 0.058 | 0.059879 | 0.0593 | 0.0594 | 0.0587 | 0.04 | 0.0595 | 0.06 | 0.064283 | 0.0533 | 0.055 | 0.065 | 0.06 |  |  |  | 1
| 1 | 2 | 0.060845 | 0.0606 | 0.059 | 0.062 | 0.059 | 0.055 | 0.060392 | 0.06 | 0.0519 | 0.0585 | 0.03 | 0.0605 | 0.06 | 0.060905 | 0.058 | 0.056 | 0.065 | 0.06 |  |  |  | 0
| 1 | 3 | 0.060204 | 0.0594 | 0.06 | 0.059 | 0.059 | 0.058 | 0.059571 | 0.0613 | 0.065 | 0.0574 | 0.03 | 0.0597 | 0.07 | 0.073274 | 0.0591 | 0.054 | 0.065 | 0.055 |  |  |  | 16
| 1 | 4 | 0.060787 | 0.0596 | 0.062 | 0.06 | 0.059 | 0.06 | 0.0588 | 0.0616 | 0.0716 | 0.0575 | 0.03 | 0.0603 | 0.06 | 0.066295 | 0.0603 | 0.055 | 0.065 | 0.06 |  |  |  | 0.05962
| 1 | 5 | 0.060165 | 0.0564 | 0.059 | 0.059 | 0.06 | 0.058 | 0.059774 | 0.0607 | 0.0711 | 0.0585 | 0.03 | 0.0607 | 0.07 | 0.061591 | 0.0623 | 0.055 | 0.06 | 0.06 |  |  |  | 41
| 1 | 6 | 0.060247 | 0.0566 | 0.061 | 0.063 | 0.058 | 0.059 | 0.061234 | 0.0594 | 0.0712 | 0.0582 | 0.03 | 0.0613 | 0.05 | 0.059261 | 0.0595 | 0.056 | 0.065 | 0.06
| Mean |  | 0.060478 | 0.058683 | 0.059833 | 0.060333 | 0.059333 | 0.058 | 0.059942 | 0.060383 | 0.065033 | 0.058133 | 0.031667 | 0.060333 | 0.061667 | 0.064268 | 0.05875 | 0.055167 | 0.064167 | 0.059167
| Median |  | 0.060433 | 0.05945 | 0.0595 | 0.0595 | 0.059 | 0.058 | 0.059826 | 0.06035 | 0.06805 | 0.05835 | 0.03 | 0.0604 | 0.06 | 0.062937 | 0.0593 | 0.055 | 0.065 | 0.06
| Std Dev. |  | 0.0003084 | 0.0017463 | 0.001472 | 0.0017512 | 0.0010328 | 0.0016733 | 0.00081816 | 0.0009704 | 0.0080271 | 0.00055377 | 0.0040825 | 0.00066232 | 0.0075277 | 0.0050746 | 0.0030317 | 0.00075277 | 0.0020412 | 0.0020412
| Rel.Std.Dev. |  | 0.5099% | 2.976% | 2.46% | 2.903% | 1.741% | 2.885% | 1.365% | 1.607% | 12.34% | 0.9526% | 12.89% | 1.098% | 12.21% | 7.896% | 5.16% | 1.365% | 3.181% | 3.45%
| PDM3 |  | 1.439% | -1.571% | 0.3579% | 1.197% | -0.4808% | -2.717% | 0.5397% | 1.28% | 9.08% | -2.494% | -46.89% | 1.197% | 3.433% | 7.796% | -1.459% | -7.469% | 7.626% | -0.7603%
| Z-Score (Absolute) |  | 0.6 | 1.24 | 0.06 | 0.46 | 0.57 | 1.33 | 0.05 | 0.51 | 5.29 | 1.81 | 29.03 | 0.46 | 0.11 | 4.5 | 0.05 | 4.86 | 4.4 | 0.74
| Analytical results for Pb in OREAS 600 (Certified Value 193 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 204.6 | 168.5 | 206 | 197 | 213 | 191 | 201.5375 | 160 | 190 | 205.3 | 195.3816 | 220 | 192 | 177 | 201 | 164 | 169.47 | 121 | 192 | 181 | 206 | 1
| 1 | 2 | 203.3954 | 166.7 | 207 | 199 | 200 | 192 | 185.6957 | 160 | 194 | 208.1 | 195.5696 | 210 | 190 | 172 | 183 | 172 | 170.47 | 123 | 194 | 188 | 210 | 0
| 1 | 3 | 202.9446 | 171.1 | 208 | 191 | 191 | 198 | 198.4951 | 150 | 194 | 212.4 | 201.3476 | 193 | 185 | 172 | 188 | 171 | 169.59 | 133 | 187 | 192 | 214 | 16
| 1 | 4 | 207.5533 | 179.2 | 215 | 194 | 199 | 192 | 187.2635 | 140 | 193 | 214.1 | 211.5569 | 188 | 190 | 182 | 181 | 168 | 167.07 | 139 | 195 | 187 | 210 | 193.085
| 1 | 5 | 203.7567 | 177.8 | 202 | 183 | 204 | 197 | 200.0768 | 150 | 193 | 208 | 203.9159 | 182 | 185 | 182 | 206 | 176 | 166.49 | 133 | 198 | 193 | 208 | 42
| 1 | 6 | 206.8667 | 189.5 | 200 | 195 | 193 | 198 | 192.6479 | 140 | 198 | 208.4 | 206.1126 | 219 | 188 | 176 | 200 | 175 | 167.98 | 141 | 195 | 196 | 214
| Mean |  | 204.8528 | 175.4667 | 206.3333 | 193.1667 | 200 | 194.6667 | 194.2861 | 150 | 193.6667 | 209.3833 | 202.314 | 202 | 188.3333 | 176.8333 | 193.1667 | 171 | 168.5117 | 131.6667 | 193.5 | 189.5 | 210.3333
| Median |  | 204.1783 | 174.45 | 206.5 | 194.5 | 199.5 | 194.5 | 195.5715 | 150 | 193.5 | 208.25 | 202.6318 | 201.5 | 189 | 176.5 | 194 | 171.5 | 168.725 | 133 | 194.5 | 190 | 210
| Std Dev. |  | 1.9172 | 8.4944 | 5.2409 | 5.6716 | 7.9498 | 3.3267 | 6.7768 | 8.9443 | 2.582 | 3.2419 | 6.2754 | 16.456 | 2.8752 | 4.4907 | 10.496 | 4.4721 | 1.5725 | 8.165 | 3.7283 | 5.3198 | 3.2042
| Rel.Std.Dev. |  | 0.9359% | 4.841% | 2.54% | 2.936% | 3.975% | 1.709% | 3.488% | 5.963% | 1.333% | 1.548% | 3.102% | 8.147% | 1.527% | 2.54% | 5.434% | 2.615% | 0.9332% | 6.201% | 1.927% | 2.807% | 1.523%
| PDM3 |  | 6.095% | -9.125% | 6.861% | 0.04231% | 3.581% | 0.8192% | 0.6221% | -22.31% | 0.3013% | 8.441% | 4.78% | 4.617% | -2.461% | -8.417% | 0.04231% | -11.44% | -12.73% | -31.81% | 0.2149% | -1.857% | 8.933%
| Z-Score (Absolute) |  | 0.87 | 1.42 | 0.99 | 0.04 | 0.49 | 0.08 | 0.05 | 3.41 | 0 | 1.23 | 0.67 | 0.65 | 0.42 | 1.31 | 0.04 | 1.77 | 1.96 | 4.83 | 0.09 | 0.32 | 1.3
| Analytical results for Pr in OREAS 600 (Certified Value 5.42 ppm) | [22 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS | 2
3
| 1 | 1 | 01: 5.14; 02: 5.2767; 05: 6.4987; 06: 5.46; 07: 5.75; 08: 5.22 | 1
| 1 | 2 | 01: 5; 02: 5.2279; 05: 6.3539; 06: 5.62; 07: 6.04; 08: 5.2 | 0
| 1 | 3 | 01: 5.12; 02: 5.2972; 05: 6.5456; 06: 5.7; 07: 5.9; 08: 5.16 | 16
| 1 | 4 | 01: 5.2; 02: 5.1374; 05: 6.1881; 06: 5.56; 07: 5.99; 08: 5.14 | 5.4201
| 1 | 5 | 01: 5.15; 02: 5.2922; 05: 6.5115; 06: 5.59; 07: 5.95; 08: 5.14 | 43
| 1 | 6 | 01: 5.05; 02: 5.1405; 05: 6.425; 06: 5.81; 07: 6.12; 08: 5.22
| Mean |  | 01: 5.11; 02: 5.2286; 05: 6.4204; 06: 5.6233; 07: 5.9583; 08: 5.18
| Median |  | 01: 5.13; 02: 5.2523; 05: 6.4618; 06: 5.605; 07: 5.97; 08: 5.18
| Std Dev. |  | 01: 0.072664; 02: 0.073674; 05: 0.13303; 06: 0.12044; 07: 0.12703; 08: 0.037947
| Rel.Std.Dev. |  | 01: 1.422%; 02: 1.409%; 05: 2.072%; 06: 2.142%; 07: 2.132%; 08: 0.7326%
| PDM3 |  | 01: -5.721%; 02: -3.532%; 05: 18.46%; 06: 3.75%; 07: 9.931%; 08: -4.429%
| Z-Score (Absolute) |  | 01: 0.76; 02: 0.47; 05: 2.39; 06: 0.47; 07: 1.28; 08: 0.59
| Analytical results for Rb in OREAS 600 (Certified Value 75 ppm) | [22 empty] | Certified
Rb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 24; 12: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS | 0
1
| 1 | 1 | 01: 73.31; 02: 74.5974; 05: 83.5; 06: 77.8757; 07: 72.1; 08: 79.08; 09: 86; 10: 68.234; 11: 66.2; 12: 77.2 | 1
| 1 | 2 | 01: 70.52; 02: 74.4214; 05: 62.8; 06: 78.0843; 07: 71.9; 08: 74.68; 09: 78.1; 10: 69.847; 11: 68.2; 12: 77.2 | 7
| 1 | 3 | 01: 73.41; 02: 74.8329; 05: 74.7; 06: 77.7619; 07: 67.8; 08: 79.46; 09: 83.1; 10: 68.15; 11: 69.6; 12: 77 | 16
| 1 | 4 | 01: 71.85; 02: 75.7718; 05: 81.9; 06: 77.3917; 07: 72.7; 08: 75.49; 09: 88.2; 10: 67.868; 11: 68.4; 12: 78.8 | 75.2896
| 1 | 5 | 01: 70.84; 02: 75.2445; 05: 77; 06: 77.747; 07: 71; 08: 77.09; 09: 90.9; 10: 69.234; 11: 68.6; 12: 77.8 | 44
| 1 | 6 | 01: 70.71; 02: 74.1707; 05: 78.7; 06: 76.7006; 07: 73.5; 08: 77.88; 09: 86.9; 10: 67.726; 11: 69.2; 12: 71.2
| Mean |  | 01: 71.7733; 02: 74.8398; 05: 76.4333; 06: 77.5935; 07: 71.5; 08: 77.28; 09: 85.5333; 10: 68.5098; 11: 68.3667; 12: 76.5333
| Median |  | 01: 71.345; 02: 74.7151; 05: 77.85; 06: 77.7544; 07: 72; 08: 77.485; 09: 86.45; 10: 68.192; 11: 68.5; 12: 77.2
| Std Dev. |  | 01: 1.3129; 02: 0.58549; 05: 7.4048; 06: 0.49204; 07: 1.995; 08: 1.9159; 09: 4.4536; 10: 0.84194; 11: 1.1827; 12: 2.6942
| Rel.Std.Dev. |  | 01: 1.829%; 02: 0.7823%; 05: 9.688%; 06: 0.6341%; 07: 2.79%; 08: 2.479%; 09: 5.207%; 10: 1.229%; 11: 1.73%; 12: 3.52%
| PDM3 |  | 01: -4.67%; 02: -0.5975%; 05: 1.519%; 06: 3.06%; 07: -5.033%; 08: 2.644%; 09: 13.61%; 10: -9.005%; 11: -9.195%; 12: 1.652%
| Z-Score (Absolute) |  | 01: 0.84; 02: 0.24; 05: 0.6; 06: 0.3; 07: 0.74; 08: 0.24; 09: 1.85; 10: 1.47; 11: 1.5; 12: 0.3
| Analytical results for Re in OREAS 600 (Indicative Value < 2 ppb) | [22 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 14; 08: 15; 09: 25 | ppb
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS | 0
0
| 1 | 1 | 01: < 2; 02: 1.931; 05: < 5000; 06: 3; 07: < 50; 08: 10; 09: < 100 | 1
| 1 | 2 | 01: < 2; 02: < 2; 05: < 5000; 06: 3; 07: < 50; 08: < 10; 09: < 100 | 8
| 1 | 3 | 01: < 2; 02: < 2; 05: < 5000; 06: 3; 07: < 50; 08: < 10; 09: < 100 | 16
| 1 | 4 | 01: < 2; 02: 1.8858; 05: < 5000; 06: 2; 07: < 50; 08: 10; 09: < 100 | < 2
| 1 | 5 | 01: < 2; 02: < 2; 05: < 5000; 06: 2; 07: < 50; 08: < 10; 09: < 100 | 9
| 1 | 6 | 01: < 2; 02: < 2; 05: < 5000; 06: 3; 07: < 50; 08: 20; 09: < 100
| Mean |  | 02: 1.9084; 06: 2.6667; 08: 13.3333
| Median |  | 02: 1.9084; 06: 3; 08: 10
| Std Dev. |  | 02: 0.031935; 06: 0.5164; 08: 5.7735
| Rel.Std.Dev. |  | 02: 1.673%; 06: 19.36%; 08: 43.3%
PDM3
| Z-Score (Absolute) |  | 01: NA; 02: 0.67; 05: NA; 06: 0; 07: NA; 08: 1.96; 09: NA
| Analytical results for S in OREAS 600 (Certified Value 1.69 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | PF*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | PF*OES |  |  | 2
3
| 1 | 1 | 1.911 | 1.6413 | 1.71 | 1.72 | 1.73 | 1.66 | 1.5941 | 1.65 | 1.61 | 2.01 | 1.6677 | 1.737 | 1.675 | N.A. | - | 1.754 | 1.6 | 1.74 | 1.64 |  |  | 1
| 1 | 2 | 1.8921 | 1.6598 | 1.72 | 1.79 | 1.69 | 1.58 | 1.6198 | 1.67 | 1.6 | 1.82 | 1.6924 | 1.7097 | 1.682 | N.A. | - | 1.679 | 1.6 | 1.73 | 1.64 |  |  | 0
| 1 | 3 | 1.8648 | 1.6301 | 1.71 | 1.67 | 1.67 | 1.67 | 1.6361 | 1.7 | 1.77 | 2.01 | 1.6647 | 1.6749 | 1.689 | N.A. | - | 1.744 | 1.6 | 1.73 | 1.68 |  |  | 16
| 1 | 4 | 1.8816 | 1.6837 | 1.77 | 1.7 | 1.7 | 1.68 | 1.5939 | 1.73 | 1.66 | 2.15 | 1.6913 | 1.6691 | 1.664 | N.A. | - | 1.671 | 1.6 | 1.72 | 1.68 |  |  | 1.691
| 1 | 5 | 1.8438 | 1.6675 | 1.68 | 1.68 | 1.67 | 1.65 | 1.6205 | 1.74 | 1.71 | 2.14 | 1.6755 | 1.6646 | 1.665 | N.A. | - | 1.706 | 1.6 | 1.68 | 1.64 |  |  | 45
| 1 | 6 | 1.9635 | 1.6597 | 1.77 | 1.75 | 1.66 | 1.69 | 1.5623 | 1.77 | 1.69 | 2.06 | 1.6735 | 1.6964 | 1.664 | N.A. | - | 1.688 | 1.6 | 1.75 | 1.72
| Mean |  | 1.8928 | 1.657 | 1.7267 | 1.7183 | 1.6867 | 1.655 | 1.6044 | 1.71 | 1.6733 | 2.0317 | 1.6775 | 1.692 | 1.6732 |  |  | 1.707 | 1.6 | 1.725 | 1.6667
| Median |  | 1.8868 | 1.6598 | 1.715 | 1.71 | 1.68 | 1.665 | 1.6069 | 1.715 | 1.675 | 2.035 | 1.6745 | 1.6857 | 1.67 |  |  | 1.697 | 1.6 | 1.73 | 1.66
| Std Dev. |  | 0.041571 | 0.019018 | 0.036148 | 0.04535 | 0.02582 | 0.03937 | 0.026409 | 0.045166 | 0.064083 | 0.12024 | 0.011785 | 0.028019 | 0.010647 |  |  | 0.034699 | 2.4324e-16 | 0.02429 | 0.03266
| Rel.Std.Dev. |  | 2.196% | 1.148% | 2.094% | 2.639% | 1.531% | 2.379% | 1.646% | 2.641% | 3.83% | 5.918% | 0.7025% | 1.656% | 0.6364% |  |  | 2.033% | 1.52e-14% | 1.408% | 1.96%
| PDM3 |  | 11.93% | -2.011% | 2.108% | 1.615% | -0.2577% | -2.13% | -5.12% | 1.122% | -1.046% | 20.14% | -0.7989% | 0.05474% | -1.056% |  |  | 0.9447% | -5.383% | 2.009% | -1.44%
| Z-Score (Absolute) |  | 5.96 | 0.86 | 1.16 | 0.92 | 0 | 0.48 | 2.38 | 0.67 | 0.39 | 9.97 | 0.26 | 0.15 | 0.39 | NA | NA | 0.59 | 2.5 | 1.11 | 0.58
| Analytical results for Sb in OREAS 600 (Certified Value 14.3 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  | 1
1
| 1 | 1 | 13.5867 | 14.18 | 17 | 14 | 15 | 15 | 14.0296 | 20 | 9 | 10.67 | 9.3807 | 14.2 | 14.52 | 14 | 13.7 | 13 | 16.137 | 11 | 17 | 15.4 |  | 1
| 1 | 2 | 13.7333 | 13.78 | 14 | 17 | 15 | 11 | 13.8643 | 20 | 10 | 9.15 | 9.5256 | 15.1 | 15.02 | 14 | 15.3 | 15 | 16.107 | 11 | 11 | 16.2 |  | 0
| 1 | 3 | 13.5033 | 14.26 | 16 | 13 | 18 | 15 | 13.9407 | 20 | 12 | 8.18 | 9.748 | 14.1 | 14.76 | 12 | 16.6 | 11 | 16.105 | 15 | 13 | 14.6 |  | 16
| 1 | 4 | 13.8067 | 14.29 | 15 | 15 | 15 | 11 | 13.9081 | < 10 | 13 | 10.38 | 9.269 | 13.7 | 15.51 | 12 | 15.5 | 13 | 16.622 | 17 | 12 | 14.8 |  | 14.3374
| 1 | 5 | 13.8633 | 13.64 | 14 | 16 | 17 | 14 | 14.1705 | 20 | 9 | 9.31 | 9.3354 | 14.7 | 14.98 | 16 | 15.9 | 13 | 15.504 | 15 | 15 | 16.4 |  | 46
| 1 | 6 | 13.7 | 14.22 | 17 | 16 | 19 | 14 | 13.6993 | 20 | 12 | 9.58 | 9.3483 | 14.4 | 14.29 | 14 | 15.6 | 13 | 16.607 | 13 | 14 | 15.8
| Mean |  | 13.6989 | 14.0617 | 15.5 | 15.1667 | 16.5 | 13.3333 | 13.9354 | 20 | 10.8333 | 9.545 | 9.4345 | 14.3667 | 14.8467 | 13.6667 | 15.4333 | 13 | 16.1803 | 13.6667 | 13.6667 | 15.5333
| Median |  | 13.7167 | 14.2 | 15.5 | 15.5 | 16 | 14 | 13.9244 | 20 | 11 | 9.445 | 9.3645 | 14.3 | 14.87 | 14 | 15.55 | 13 | 16.122 | 14 | 13.5 | 15.6
| Std Dev. |  | 0.13467 | 0.27845 | 1.3784 | 1.472 | 1.7607 | 1.8619 | 0.15855 | 0 | 1.7224 | 0.89873 | 0.17556 | 0.48854 | 0.42707 | 1.5055 | 0.96264 | 1.2649 | 0.4117 | 2.4221 | 2.1602 | 0.73394
| Rel.Std.Dev. |  | 0.9831% | 1.98% | 8.893% | 9.705% | 10.67% | 13.96% | 1.138% | 0% | 15.9% | 9.416% | 1.861% | 3.4% | 2.877% | 11.02% | 6.237% | 9.73% | 2.544% | 17.72% | 15.81% | 4.725%
| PDM3 |  | -4.454% | -1.923% | 8.109% | 5.784% | 15.08% | -7.003% | -2.804% | 39.5% | -24.44% | -33.43% | -34.2% | 0.2039% | 3.552% | -4.678% | 7.644% | -9.328% | 12.85% | -4.678% | -4.678% | 8.341%
| Z-Score (Absolute) |  | 0.19 | 0.04 | 0.93 | 0.73 | 1.56 | 0.41 | 0.04 | 3.74 | 1.97 | 2.77 | 2.88 | 0.23 | 0.53 | 0.21 | 1.11 | 0.62 | 1.36 | 0.21 | 0.21 | 0.96
| Analytical results for Sc in OREAS 600 (Certified Value 5.95 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 5.7805 | 5.7 | 6 | 6 | 6 | 6 | 5.5918 | 6 | 5.4 | 6.0751 | 6 | 6 | 4.3 | 7.6 | 6 | 6.029 | 5 | 6 | 6 |  |  | 1
| 1 | 2 | 5.8913 | 5.4 | 6 | 6 | 6 | 5 | 5.8923 | 6 | 5.1 | 5.8451 | 6 | 6 | 4.2 | 8.1 | 6 | 6.603 | 7 | 6 | 6 |  |  | 0
| 1 | 3 | 5.5312 | 5.5 | 6 | 6 | 6 | 6 | 5.6998 | 6 | 5.1 | 5.8477 | 6 | 6 | 4.4 | 7.7 | 6 | 6.588 | 7 | 6 | 6 |  |  | 16
| 1 | 4 | 5.7053 | 5.3 | 6 | 6 | 6 | 6 | 6.0799 | 5 | 5.4 | 5.9171 | 6 | 6 | 4.2 | 7.9 | 6 | 6.464 | 7 | 6 | 6 |  |  | 5.9507
| 1 | 5 | 5.5418 | 5.3 | 6 | 6 | 6 | 6 | 5.7819 | 6 | 5.3 | 5.896 | 6 | 6 | 4.5 | 8 | 6 | 6.43 | 7 | 6 | 6 |  |  | 47
| 1 | 6 | 6.1482 | 5.3 | 6 | 6 | 6 | 6 | 5.8022 | 6 | 5.1 | 5.7445 | 6 | 6 | 4.7 | 8.1 | 6 | 6.173 | 7 | 6 | 6
| Mean |  | 5.7664 | 5.4167 | 6 | 6 | 6 | 5.8333 | 5.808 | 5.8333 | 5.2333 | 5.8876 | 6 | 6 | 4.3833 | 7.9 | 6 | 6.3812 | 6.6667 | 6 | 6
| Median |  | 5.7429 | 5.35 | 6 | 6 | 6 | 6 | 5.792 | 6 | 5.2 | 5.8718 | 6 | 6 | 4.35 | 7.95 | 6 | 6.447 | 7 | 6 | 6
| Std Dev. |  | 0.23285 | 0.16021 | 0 | 0 | 0 | 0.40825 | 0.16746 | 0.40825 | 0.15055 | 0.10952 | 0 | 0 | 0.19408 | 0.20976 | 0 | 0.23177 | 0.8165 | 0 | 0
| Rel.Std.Dev. |  | 4.038% | 2.958% | 0% | 0% | 0% | 6.999% | 2.883% | 6.999% | 2.877% | 1.86% | 0% | 0% | 4.428% | 2.655% | 0% | 3.632% | 12.25% | 0% | 0%
| PDM3 |  | -3.097% | -8.974% | 0.8293% | 0.8293% | 0.8293% | -1.972% | -2.398% | -1.972% | -12.05% | -1.06% | 0.8293% | 0.8293% | -26.34% | 32.76% | 0.8293% | 7.235% | 12.03% | 0.8293% | 0.8293%
Z-Score (Absolute) | bar: 01 to 25
| Analytical results for Se in OREAS 600 (Certified Value 6.97 ppm) | [22 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 21 | 24 | 25 | [8 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [8 empty] | 2
2
| 1 | 1 | 7.6428 | 6.5 | 6.5467 | 6 | 8 | 7 | 7.7 | 2 | 29 | 5.855 | 30 | 10 | < 5 | [8 empty] | 1
| 1 | 2 | 7.6998 | 6.7 | 6.991 | 7 | 6 | 7 | 6.65 | < 2 | 30 | 5.525 | 25 | 10 | 5 | [8 empty] | 0
| 1 | 3 | 8.1523 | 7.2 | 6.9385 | 5 | 8 | 7 | 6.19 | < 2 | 36 | 5.352 | 38 | 10 | 5 | [8 empty] | 16
| 1 | 4 | 8.038 | 6.9 | 6.5294 | 7 | 8 | 7 | 6.83 | < 2 | 40 | 5.88 | 32 | 10 | 5 | [8 empty] | 6.974
| 1 | 5 | 7.521 | 6.6 | 6.8139 | 7 | 8 | 7 | 5.92 | < 2 | 34 | 5.438 | 34 | 10 | 5 | [8 empty] | 48
| 1 | 6 | 8.095 | 7.1 | 6.9236 | 6 | 8 | 7 | 6.62 | < 2 | 40 | 5.588 | 35 | 10 | 5
| Mean |  | 7.8582 | 6.8333 | 6.7905 | 6.3333 | 7.6667 | 7 | 6.6517 | 2 | 34.8333 | 5.6063 | 32.3333 | 10 | 5
| Median |  | 7.8689 | 6.8 | 6.8688 | 6.5 | 8 | 7 | 6.635 | 2 | 35 | 5.5565 | 33 | 10 | 5
| Std Dev. |  | 0.26836 | 0.28048 | 0.20396 | 0.8165 | 0.8165 | 0 | 0.61323 |  | 4.7504 | 0.21758 | 4.5019 | 0 | 0
| Rel.Std.Dev. |  | 3.415% | 4.105% | 3.004% | 12.89% | 10.65% | 0% | 9.219% |  | 13.64% | 3.881% | 13.92% | 0% | 0%
| PDM3 |  | 12.68% | -2.017% | -2.631% | -9.187% | 9.932% | 0.3728% | -4.622% | -71.32% | 399.5% | -19.61% | 363.6% | 43.39% | -28.31%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.03 | 0.33 | 0.55 | 0.11 | 0.12 | 3.18 | 18.42 | 0.81 | 16.78 | 2.08 | 1.21
| Analytical results for Sm in OREAS 600 (Certified Value 3.56 ppm) | [22 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  |  | 2
3
| 1 | 1 | 01: 3.39; 02: 3.8287; 05: 3.9797; 06: 3.55; 07: 3.5; 08: 3.35 |  |  | 1
| 1 | 2 | 01: 3.44; 02: 3.679; 05: 3.8332; 06: 3.73; 07: 3.1; 08: 3.35 |  |  | 0
| 1 | 3 | 01: 3.36; 02: 3.6139; 05: 3.9267; 06: 3.66; 07: 3.3; 08: 3.15 |  |  | 16
| 1 | 4 | 01: 3.41; 02: 3.5745; 05: 3.8945; 06: 3.56; 07: 3.4; 08: 3.25 |  |  | 3.5594
| 1 | 5 | 01: 3.54; 02: 3.789; 05: 3.8961; 06: 3.68; 07: 3.5; 08: 3.45 |  |  | 49
| 1 | 6 | 01: 3.29; 02: 3.5515; 05: 3.9301; 06: 3.73; 07: 3.4; 08: 3.35
| Mean |  | 01: 3.405; 02: 3.6727; 05: 3.9101; 06: 3.6517; 07: 3.3667; 08: 3.3167
| Median |  | 01: 3.4; 02: 3.6464; 05: 3.9114; 06: 3.67; 07: 3.4; 08: 3.35
| Std Dev. |  | 01: 0.083606; 02: 0.11461; 05: 0.048724; 06: 0.079854; 07: 0.15055; 08: 0.10328
| Rel.Std.Dev. |  | 01: 2.455%; 02: 3.121%; 05: 1.246%; 06: 2.187%; 07: 4.472%; 08: 3.114%
| PDM3 |  | 01: -4.337%; 02: 3.186%; 05: 9.853%; 06: 2.593%; 07: -5.414%; 08: -6.818%
| Z-Score (Absolute) |  | 01: 0.54; 02: 0.64; 05: 1.68; 06: 0.54; 07: 0.71; 08: 0.79
| Analytical results for Sn in OREAS 600 (Certified Value 2.12 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [7 empty] | 2
3
| 1 | 1 | < 5 | 2.3 | 2.0745 | < 10 | 2 | 7.6 | 2.38 | < 15 | 1.8 | < 20 | 1.931 | 3 | 2.2 | 2 | [7 empty] | 1
| 1 | 2 | < 5 | 2.4 | 2.0949 | < 10 | 1.6 | 5.2 | 2.52 | < 15 | 1.9 | < 20 | 2.037 | 3 | 2.3 | 2 | [7 empty] | 0
| 1 | 3 | < 5 | 2.3 | 2.0276 | < 10 | 2.1 | 2.2 | 2.46 | < 15 | 1.8 | < 20 | 1.909 | < 2 | 2.2 | 2 | [7 empty] | 16
| 1 | 4 | < 5 | 2.3 | 2.2068 | < 10 | 2.2 | 2.2 | 2.42 | < 15 | 1.8 | < 20 | 1.957 | 3 | 2.1 | 2 | [7 empty] | 2.117
| 1 | 5 | < 5 | 2.3 | 2.1177 | < 10 | 2.2 | 2.2 | 2.28 | < 15 | 1.8 | < 20 | 1.928 | 3 | 2.3 | 2 | [7 empty] | 50
| 1 | 6 | < 5 | 2.3 | 2.0631 | < 10 | 2.1 | 5.2 | 2.27 | < 15 | 1.8 | < 20 | 1.918 | 3 | 2.4 | 2
| Mean |  | 02: 2.3167; 05: 2.0974; 07: 2.0333; 08: 4.1; 09: 2.3883; 12: 1.8167; 14: 1.9467; 15: 3; 16: 2.25; 17: 2
| Median |  | 02: 2.3; 05: 2.0847; 07: 2.1; 08: 3.7; 09: 2.4; 12: 1.8; 14: 1.9295; 15: 3; 16: 2.25; 17: 2
| Std Dev. |  | 02: 0.040825; 05: 0.061601; 07: 0.22509; 08: 2.2583; 09: 0.099281; 12: 0.040825; 14: 0.04712; 15: 0; 16: 0.10488; 17: 0
| Rel.Std.Dev. |  | 02: 1.762%; 05: 2.937%; 07: 11.07%; 08: 55.08%; 09: 4.157%; 12: 2.247%; 14: 2.421%; 15: 0%; 16: 4.661%; 17: 0%
| PDM3 |  | 02: 9.433%; 05: -0.923%; 07: -3.951%; 08: 93.67%; 09: 12.82%; 12: -14.19%; 14: -8.045%; 15: 41.71%; 16: 6.284%; 17: -5.525%
| Z-Score (Absolute) |  | NA | 0.46 | 0.3 | NA | 0.23 | 6.65 | 0.71 | NA | 1.28 | NA | 0.83 | 2.83 | 0.23 | 0.64
| Analytical results for Sr in OREAS 600 (Certified Value 186 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | PF*MS |  | 0
0
| 1 | 1 | 204.29 | 190.36 | 185 | 182 | 193 | 186 | 187.1323 | 178 | 210 | 193.8862 | 179.1 | 182 | 121 | 169 | 178 | 183.23 | 164 | 161 | 184 | 200 |  | 1
| 1 | 2 | 197.81 | 184.75 | 187 | 186 | 188 | 178 | 186.9012 | 182 | 182 | 194.5752 | 178.3 | 188 | 115 | 171 | 187 | 178.78 | 180 | 164 | 190 | 190 |  | 0
| 1 | 3 | 197.705 | 190.6 | 187 | 177 | 185 | 189 | 188.6472 | 184 | 198.4 | 194.8152 | 177.3 | 187 | 124 | 175 | 183 | 175.14 | 185 | 163 | 189 | 190 |  | 16
| 1 | 4 | 200.77 | 185.73 | 194 | 178 | 189 | 189 | 188.8655 | 185 | 204.9 | 194.4856 | 177.4 | 186 | 122 | 172 | 184 | 171.43 | 188 | 162 | 188 | 190 |  | 186.3569
| 1 | 5 | 191.77 | 181.11 | 184 | 176 | 188 | 187 | 186.8656 | 185 | 205.9 | 205.621 | 177.2 | 186 | 132 | 172 | 187 | 175.1 | 193 | 163 | 193 | 190 |  | 51
| 1 | 6 | 203.655 | 183.16 | 192 | 183 | 185 | 189 | 187.3444 | 185 | 197.4 | 196.0801 | 180.7 | 187 | 125 | 170 | 185 | 167.189 | 194 | 170 | 198 | 190
| Mean |  | 199.3333 | 185.9517 | 188.1667 | 180.3333 | 188 | 186.3333 | 187.626 | 183.1667 | 199.7667 | 196.5772 | 178.3333 | 186 | 123.1667 | 171.5 | 184 | 175.1448 | 184 | 163.8333 | 190.3333 | 191.6667
| Median |  | 199.29 | 185.24 | 187 | 180 | 188 | 188 | 187.2383 | 184.5 | 201.65 | 194.6952 | 177.85 | 186.5 | 123 | 171.5 | 184.5 | 175.12 | 186.5 | 163 | 189.5 | 190
| Std Dev. |  | 4.6365 | 3.8411 | 3.9707 | 3.9328 | 2.9665 | 4.274 | 0.89509 | 2.7869 | 9.9178 | 4.4892 | 1.3721 | 2.0976 | 5.5648 | 2.0736 | 3.3466 | 5.58 | 11.0815 | 3.1885 | 4.761 | 4.0825
| Rel.Std.Dev. |  | 2.326% | 2.066% | 2.11% | 2.181% | 1.578% | 2.294% | 0.4771% | 1.521% | 4.965% | 2.284% | 0.7694% | 1.128% | 4.518% | 1.209% | 1.819% | 3.186% | 6.023% | 1.946% | 2.501% | 2.13%
| PDM3 |  | 6.963% | -0.2175% | 0.9711% | -3.232% | 0.8817% | -0.01265% | 0.681% | -1.712% | 7.196% | 5.484% | -4.305% | -0.1915% | -33.91% | -7.972% | -1.265% | -6.016% | -1.265% | -12.09% | 2.134% | 2.849%
| Z-Score (Absolute) |  | 2.15 | 0 | 0.35 | 0.91 | 0.33 | 0.33 | 0.27 | 0.45 | 2.22 | 1.41 | 1.23 | 0 | 10.11 | 2.33 | 0.32 | 1.74 | 0.32 | 3.56 | 0.7 | 0.65
| Analytical results for Ta in OREAS 600 (Certified Value 0.63 ppm) | [22 empty] | Certified
Ta | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 14; 07: 15; 08: 17; 09: 19; 11: 24; 12: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS |  | 2
2
| 1 | 1 | 01: 1.15; 02: 0.68193; 05: 0.46; 06: 0.63; 07: 1.4; 08: 0.33; 09: 0.781; 11: 0.8; 12: 0.5 |  | 1
| 1 | 2 | 01: 0.91; 02: 0.71178; 05: 0.33; 06: 0.62; 07: 1.29; 08: 0.3; 09: 0.716; 11: 0.8; 12: 0.5 |  | 7
| 1 | 3 | 01: 0.93; 02: 0.68219; 05: 0.53; 06: 0.62; 07: 1.19; 08: 0.41; 09: 0.712; 11: 0.8; 12: 0.5 |  | 16
| 1 | 4 | 01: 0.98; 02: 0.64872; 05: 0.59; 06: 0.6; 07: 1.18; 08: 0.41; 09: 0.76; 11: 0.8; 12: 0.5 |  | 0.63323
| 1 | 5 | 01: 0.92; 02: 0.65034; 05: 0.61; 06: 0.61; 07: 1.16; 08: 0.35; 09: 0.749; 11: 0.7; 12: 0.5 |  | 52
| 1 | 6 | 01: 0.93; 02: 0.62816; 05: 0.58; 06: 0.62; 07: 1.04; 08: 0.45; 09: 0.775; 11: 0.6; 12: 0.5
| Mean |  | 01: 0.97; 02: 0.66719; 05: 0.51667; 06: 0.61667; 07: 1.21; 08: 0.375; 09: 0.74883; 11: 0.75; 12: 0.5
| Median |  | 01: 0.93; 02: 0.66614; 05: 0.555; 06: 0.62; 07: 1.185; 08: 0.38; 09: 0.7545; 11: 0.8; 12: 0.5
| Std Dev. |  | 01: 0.091433; 02: 0.030261; 05: 0.10614; 06: 0.010328; 07: 0.12264; 08: 0.057184; 09: 0.029253; 11: 0.083666; 12: 0
| Rel.Std.Dev. |  | 01: 9.426%; 02: 4.536%; 05: 20.54%; 06: 1.675%; 07: 10.14%; 08: 15.25%; 09: 3.907%; 11: 11.16%; 12: 0%
| PDM3 |  | 01: 53.18%; 02: 5.363%; 05: -18.41%; 06: -2.615%; 07: 91.09%; 08: -40.78%; 09: 18.26%; 11: 18.44%; 12: -21.04%
| Z-Score (Absolute) |  | 01: 1.2; 02: 0; 05: 0.67; 06: 0.23; 07: 2.43; 08: 1.31; 09: 0.37; 11: 0.37; 12: 0.75
| Analytical results for Tb in OREAS 600 (Certified Value 0.4 ppm) | [22 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 19; 09: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*OES/MS; 09: 4A*MS |  | 2
2
| 1 | 1 | 01: 0.35; 02: 0.41995; 05: 0.42194; 06: 0.6; 07: 0.41; 08: 0.345; 09: 0.32 |  | 1
| 1 | 2 | 01: 0.35; 02: 0.39143; 05: 0.42827; 06: 0.47; 07: 0.43; 08: 0.354; 09: 0.34 |  | 0
| 1 | 3 | 01: 0.35; 02: 0.40571; 05: 0.45694; 06: 0.46; 07: 0.41; 08: 0.344; 09: 0.32 |  | 16
| 1 | 4 | 01: 0.37; 02: 0.40878; 05: 0.42628; 06: 0.5; 07: 0.42; 08: 0.351; 09: 0.32 |  | 0.3967
| 1 | 5 | 01: 0.36; 02: 0.42543; 05: 0.47464; 06: 0.42; 07: 0.42; 08: 0.34; 09: 0.32 |  | 53
| 1 | 6 | 01: 0.36; 02: 0.4357; 05: 0.44014; 06: 0.51; 07: 0.43; 08: 0.34; 09: 0.34
| Mean |  | 01: 0.35667; 02: 0.4145; 05: 0.44137; 06: 0.49333; 07: 0.42; 08: 0.34567; 09: 0.32667
| Median |  | 01: 0.355; 02: 0.41436; 05: 0.43421; 06: 0.485; 07: 0.42; 08: 0.3445; 09: 0.32
| Std Dev. |  | 01: 0.008165; 02: 0.01575; 05: 0.020627; 06: 0.06121; 07: 0.0089443; 08: 0.0057504; 09: 0.010328
| Rel.Std.Dev. |  | 01: 2.289%; 02: 3.8%; 05: 4.673%; 06: 12.41%; 07: 2.13%; 08: 1.664%; 09: 3.162%
| PDM3 |  | 01: -10.09%; 02: 4.488%; 05: 11.26%; 06: 24.36%; 07: 5.875%; 08: -12.86%; 09: -17.65%
| Z-Score (Absolute) |  | 01: 0.68; 02: 0; 05: 0.32; 06: 0.67; 07: 0.06; 08: 0.81; 09: 1.03
| Analytical results for Te in OREAS 600 (Certified Value 7.34 ppm) | [22 empty] | Certified
Te | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 13; 08: 14; 09: 15; 11: 17; 12: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 11: 4A*OES/MS; 12: 4A*MS |  | 2
3
| 1 | 1 | 01: 7.2; 02: 6.8136; 05: 7; 06: 7.44; 07: 5.393; 08: 7.1; 09: 8.13; 11: 4.33; 12: 7 |  | 1
| 1 | 2 | 01: 7.3; 02: 6.8152; 05: 8; 06: 5.89; 07: 5.0001; 08: 7.2; 09: 8.61; 11: 4.9; 12: 7.2 |  | 0
| 1 | 3 | 01: 7; 02: 6.8063; 05: 7; 06: 6.7; 07: 5.2775; 08: 7.1; 09: 7.85; 11: 4.04; 12: 7.2 |  | 16
| 1 | 4 | 01: 7.8; 02: 6.8421; 05: 7; 06: 7.2; 07: 5.3192; 08: 7; 09: 8.47; 11: 4.3; 12: 6.8 |  | 7.3446
| 1 | 5 | 01: 7.6; 02: 6.8789; 05: 8; 06: 7.92; 07: 5.3738; 08: 7; 09: 8.33; 11: 4.4; 12: 6.6 |  | 54
| 1 | 6 | 01: 8.1; 02: 6.7771; 05: 8; 06: 7.52; 07: 5.0096; 08: 7.4; 09: 8.08; 11: 4.12; 12: 7.8
| Mean |  | 01: 7.5; 02: 6.8222; 05: 7.5; 06: 7.1117; 07: 5.2289; 08: 7.1333; 09: 8.245; 11: 4.3483; 12: 7.1
| Median |  | 01: 7.45; 02: 6.8144; 05: 7.5; 06: 7.32; 07: 5.2984; 08: 7.1; 09: 8.23; 11: 4.315; 12: 7.1
| Std Dev. |  | 01: 0.40988; 02: 0.03469; 05: 0.54772; 06: 0.72079; 07: 0.17826; 08: 0.15055; 09: 0.27841; 11: 0.30229; 12: 0.41473
| Rel.Std.Dev. |  | 01: 5.465%; 02: 0.5085%; 05: 7.303%; 06: 10.14%; 07: 3.409%; 08: 2.111%; 09: 3.377%; 11: 6.952%; 12: 5.841%
| PDM3 |  | 01: 2.116%; 02: -7.113%; 05: 2.116%; 06: -3.172%; 07: -28.81%; 08: -2.876%; 09: 12.26%; 11: -40.8%; 12: -3.33%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.5; 05: 0.67; 06: 0; 07: 3.27; 08: 0.04; 09: 1.97; 11: 4.99; 12: 0.02
| Analytical results for Th in OREAS 600 (Certified Value 9.38 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 8.98 | < 20 | < 20 | < 20 | < 20 | 9.6105 | 5.8 | 11.413 | 9.24 | 8.99 | 9.2 | < 20 | 8.49 | 13 | 9.5 | 9.7 | [5 empty] | 1
| 1 | 2 | 8.91 | < 20 | < 20 | < 20 | < 20 | 9.4632 | 4 | 10.9852 | 9.5 | 9.65 | 9.7 | < 20 | 8.5 | 12 | 10 | 9.9 | [5 empty] | 0
| 1 | 3 | 8.74 | < 20 | < 20 | < 20 | < 20 | 9.4497 | 7.7 | 11.4404 | 8.72 | 9.59 | 9.4 | < 20 | 8.19 | 13 | 9.9 | 10.4 | [5 empty] | 16
| 1 | 4 | 9.22 | < 20 | < 20 | < 20 | < 20 | 9.1635 | 7.5 | 11.232 | 9.22 | 9.61 | 9.6 | < 20 | 8.5 | 12 | 10 | 10.3 | [5 empty] | 9.3845
| 1 | 5 | 8.94 | < 20 | < 20 | < 20 | < 20 | 9.3935 | 7.8 | 11.0502 | 9.03 | 9.72 | 9.8 | < 20 | 8.57 | 13 | 10 | 9.9 | [5 empty] | 55
| 1 | 6 | 9.04 | < 20 | < 20 | < 20 | < 20 | 9.2761 | 7.5 | 11.1664 | 9.4 | 9.48 | 9.8 | < 20 | 8.55 | 12 | 9.7 | 9.9
| Mean |  | 8.9717 |  |  |  |  | 9.3927 | 6.7167 | 11.2145 | 9.185 | 9.5067 | 9.5833 |  | 8.4667 | 12.5 | 9.85 | 10.0167
| Median |  | 8.96 |  |  |  |  | 9.4216 | 7.5 | 11.1992 | 9.23 | 9.6 | 9.65 |  | 8.5 | 12.5 | 9.95 | 9.9
| Std Dev. |  | 0.15804 |  |  |  |  | 0.15607 | 1.5224 | 0.18581 | 0.27927 | 0.26508 | 0.24014 |  | 0.13924 | 0.54772 | 0.20736 | 0.27142
| Rel.Std.Dev. |  | 1.762% |  |  |  |  | 1.662% | 22.67% | 1.657% | 3.04% | 2.788% | 2.506% |  | 1.645% | 4.382% | 2.105% | 2.71%
| PDM3 |  | -4.399% |  |  |  |  | 0.08766% | -28.43% | 19.5% | -2.126% | 1.302% | 2.119% |  | -9.78% | 33.2% | 4.96% | 6.736%
| Z-Score (Absolute) |  | 0.95 | NA | NA | NA | NA | 0.3 | 3.62 | 2.54 | 0.62 | 0.04 | 0 | NA | 1.74 | 4.54 | 0.41 | 0.67
| Analytical results for Ti in OREAS 600 (Certified Value 0.242 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES |  | 3
3
| 1 | 1 | 0.255 | 0.2481 | 0.24 | 0.25 | 0.24 | 0.23 | 0.24941 | 0.21 | 0.219 | 0.24938 | 0.2499 | 0.24 | 0.2 | 0.23 | 0.24 | 0.226 | 0.1713 | 0.23 | 0.275 | 0.25 |  | 1
| 1 | 2 | 0.25 | 0.241 | 0.25 | 0.26 | 0.25 | 0.22 | 0.25038 | 0.21 | 0.201 | 0.25229 | 0.2479 | 0.24 | 0.2 | 0.23 | 0.25 | 0.221 | 0.1865 | 0.24 | 0.27 | 0.25 |  | 0
| 1 | 3 | 0.245 | 0.2417 | 0.25 | 0.25 | 0.24 | 0.24 | 0.25332 | 0.22 | 0.22 | 0.25283 | 0.246 | 0.24 | 0.19 | 0.23 | 0.25 | 0.22 | 0.1905 | 0.23 | 0.27 | 0.26 |  | 16
| 1 | 4 | 0.245 | 0.2442 | 0.25 | 0.25 | 0.24 | 0.24 | 0.25635 | 0.23 | 0.234 | 0.25162 | 0.2439 | 0.24 | 0.2 | 0.23 | 0.25 | 0.211 | 0.1964 | 0.24 | 0.275 | 0.25 |  | 0.24219
| 1 | 5 | 0.24 | 0.2407 | 0.24 | 0.24 | 0.25 | 0.24 | 0.25208 | 0.22 | 0.235 | 0.25263 | 0.248 | 0.24 | 0.2 | 0.23 | 0.25 | 0.216 | 0.2032 | 0.24 | 0.27 | 0.26 |  | 56
| 1 | 6 | 0.255 | 0.2396 | 0.25 | 0.25 | 0.24 | 0.24 | 0.25259 | 0.21 | 0.221 | 0.25657 | 0.2441 | 0.24 | 0.19 | 0.23 | 0.25 | 0.233 | 0.1994 | 0.24 | 0.275 | 0.26
| Mean |  | 0.24833 | 0.24255 | 0.24667 | 0.25 | 0.24333 | 0.235 | 0.25236 | 0.21667 | 0.22167 | 0.25255 | 0.24663 | 0.24 | 0.19667 | 0.23 | 0.24833 | 0.22117 | 0.19122 | 0.23667 | 0.2725 | 0.255
| Median |  | 0.2475 | 0.24135 | 0.25 | 0.25 | 0.24 | 0.24 | 0.25234 | 0.215 | 0.2205 | 0.25246 | 0.24695 | 0.24 | 0.2 | 0.23 | 0.25 | 0.2205 | 0.19345 | 0.24 | 0.2725 | 0.255
| Std Dev. |  | 0.0060553 | 0.0031233 | 0.005164 | 0.0063246 | 0.005164 | 0.0083666 | 0.0024341 | 0.008165 | 0.012388 | 0.0023314 | 0.0023847 | 0 | 0.005164 | 0 | 0.0040825 | 0.007679 | 0.011462 | 0.005164 | 0.0027386 | 0.0054772
| Rel.Std.Dev. |  | 2.438% | 1.288% | 2.094% | 2.53% | 2.122% | 3.56% | 0.9645% | 3.768% | 5.589% | 0.9231% | 0.9669% | 0% | 2.626% | 0% | 1.644% | 3.472% | 5.994% | 2.182% | 1.005% | 2.148%
| PDM3 |  | 2.536% | 0.1485% | 1.848% | 3.225% | 0.472% | -2.969% | 4.198% | -10.54% | -8.474% | 4.279% | 1.835% | -0.9044% | -18.8% | -5.033% | 2.536% | -8.681% | -21.05% | -2.281% | 12.51% | 5.289%
| Z-Score (Absolute) |  | 0.42 | 0.03 | 0.29 | 0.55 | 0.03 | 0.62 | 0.73 | 2.04 | 1.65 | 0.75 | 0.29 | 0.23 | 3.6 | 1.01 | 0.42 | 1.69 | 4.02 | 0.49 | 2.3 | 0.94
| Analytical results for Tl in OREAS 600 (Certified Value 1.11 ppm) | [22 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | [6 empty] | 2
3
| 1 | 1 | < 5 | 1.13 | < 10 | < 10 | 1.0795 | < 5 | 1.29 | 0.86777 | 1.06 | 1.09 | 1.11 | < 20 | 1.143 | 1.1 | 1 | [6 empty] | 1
| 1 | 2 | < 5 | 1.16 | < 10 | < 10 | 1.0897 | < 5 | 1.03 | 0.86687 | 1.05 | 1.12 | 1.18 | < 20 | 1.137 | 1.1 | 1.1 | [6 empty] | 0
| 1 | 3 | < 5 | 1.13 | < 10 | < 10 | 1.0809 | < 5 | 1.42 | 0.91177 | 1.06 | 1.06 | 1.11 | < 20 | 1.076 | 1.2 | 1.2 | [6 empty] | 16
| 1 | 4 | < 5 | 1.18 | < 10 | < 10 | 1.0719 | < 5 | 1.2 | 0.87543 | 1.1 | 1.11 | 1.15 | < 20 | 1.091 | 1.1 | 1.1 | [6 empty] | 1.1129
| 1 | 5 | < 5 | 1.17 | < 10 | < 10 | 1.0484 | < 5 | 1.13 | 0.84805 | 1.05 | 1.1 | 1.13 | < 20 | 1.08 | 1.2 | 1.1 | [6 empty] | 57
| 1 | 6 | < 5 | 1.08 | < 10 | < 10 | 1.073 | < 5 | 1.11 | 0.85106 | 1.07 | 1.05 | 1.15 | < 20 | 1.057 | 1.3 | 1.1
| Mean |  | 02: 1.1417; 07: 1.0739; 09: 1.1967; 11: 0.87016; 12: 1.065; 13: 1.0883; 14: 1.1383; 16: 1.0973; 17: 1.1667; 18: 1.1
| Median |  | 02: 1.145; 07: 1.0763; 09: 1.165; 11: 0.86732; 12: 1.06; 13: 1.095; 14: 1.14; 16: 1.0855; 17: 1.15; 18: 1.1
| Std Dev. |  | 02: 0.03656; 07: 0.014037; 09: 0.14024; 11: 0.022932; 12: 0.018708; 13: 0.027869; 14: 0.027142; 16: 0.034875; 17: 0.08165; 18: 0.063246
| Rel.Std.Dev. |  | 02: 3.202%; 07: 1.307%; 09: 11.72%; 11: 2.635%; 12: 1.757%; 13: 2.561%; 14: 2.384%; 16: 3.178%; 17: 6.999%; 18: 5.75%
| PDM3 |  | 02: 2.583%; 07: -3.506%; 09: 7.525%; 11: -21.81%; 12: -4.305%; 13: -2.209%; 14: 2.284%; 16: -1.4%; 17: 4.83%; 18: -1.16%
| Z-Score (Absolute) |  | NA | 0.72 | NA | NA | 0.42 | NA | 0.9 | 3.84 | 0.68 | 0.17 | 0.67 | NA | 0.02 | 1.14 | 0.02
| Analytical results for Tm in OREAS 600 (Indicative Value 0.12 ppm) | [22 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  | 2
2
| 1 | 1 | 01: 0.12; 02: 0.13055; 05: 0.86777; 06: 0.37; 07: 0.11; 08: 0.12 |  | 1
| 1 | 2 | 01: 0.12; 02: 0.13596; 05: 0.86687; 06: 0.2; 07: 0.12; 08: 0.12 |  | 4
| 1 | 3 | 01: 0.12; 02: 0.1297; 05: 0.87177; 06: 0.2; 07: 0.11; 08: 0.12 |  | 16
| 1 | 4 | 01: 0.13; 02: 0.12902; 05: 0.87543; 06: 0.22; 07: 0.12; 08: 0.12 |  | 0.12242
| 1 | 5 | 01: 0.12; 02: 0.13599; 05: 0.84805; 06: 0.21; 07: 0.12; 08: 0.12 |  | 10
| 1 | 6 | 01: 0.12; 02: 0.12678; 05: 0.89106; 06: 0.24; 07: 0.12; 08: 0.12
| Mean |  | 01: 0.12167; 02: 0.13133; 05: 0.87016; 06: 0.24; 07: 0.11667; 08: 0.12
| Median |  | 01: 0.12; 02: 0.13012; 05: 0.86977; 06: 0.215; 07: 0.12; 08: 0.12
| Std Dev. |  | 01: 0.0040825; 02: 0.0038068; 05: 0.013946; 06: 0.065422; 07: 0.005164; 08: 0
| Rel.Std.Dev. |  | 01: 3.355%; 02: 2.899%; 05: 1.603%; 06: 27.26%; 07: 4.426%; 08: 0%
| PDM3 |  | 01: -0.6123%; 02: 7.283%; 05: 610.8%; 06: 96.05%; 07: -4.697%; 08: -1.974%
| Z-Score (Absolute) |  | 01: 0.4; 02: 0.4; 05: 61.41; 06: 7.23; 07: 0.81; 08: 0.54
| Analytical results for U in OREAS 600 (Certified Value 2.69 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 07 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [7 empty] | 2
3
| 1 | 1 | 2.67 | < 10 | < 10 | 2.7588 | 3 | 3.2011 | 2.59 | 2.78 | 2.25 | < 20 | 2.44 | < 20 | 2.6 | 2.6 | [7 empty] | 1
| 1 | 2 | 2.63 | < 10 | < 10 | 2.7722 | 2.4 | 3.1972 | 2.64 | 2.88 | 2.28 | < 20 | 2.44 | < 20 | 2.7 | 2.8 | [7 empty] | 0
| 1 | 3 | 2.66 | < 10 | < 10 | 2.6658 | 2.9 | 3.2328 | 2.51 | 2.8 | 2.24 | < 20 | 2.33 | < 20 | 2.8 | 2.8 | [7 empty] | 16
| 1 | 4 | 2.7 | < 10 | < 10 | 2.6694 | 3 | 3.2863 | 2.6 | 2.82 | 2.31 | < 20 | 2.43 | < 20 | 2.7 | 2.8 | [7 empty] | 2.6866
| 1 | 5 | 2.75 | < 10 | < 10 | 2.686 | 2.8 | 3.2095 | 2.53 | 2.82 | 2.32 | < 20 | 2.4 | < 20 | 2.8 | 2.6 | [7 empty] | 58
| 1 | 6 | 2.65 | < 10 | < 10 | 2.6735 | 2.8 | 3.2126 | 2.65 | 2.86 | 2.31 | < 20 | 2.45 | < 20 | 2.9 | 2.6
| Mean |  | 01: 2.6767; 06: 2.7043; 07: 2.8167; 08: 3.2232; 09: 2.5867; 11: 2.8267; 12: 2.285; 14: 2.415; 16: 2.75; 17: 2.7
| Median |  | 01: 2.665; 06: 2.6797; 07: 2.85; 08: 3.211; 09: 2.595; 11: 2.82; 12: 2.295; 14: 2.435; 16: 2.75; 17: 2.7
| Std Dev. |  | 01: 0.04274; 06: 0.048107; 07: 0.22286; 08: 0.033294; 09: 0.056804; 11: 0.037238; 12: 0.033912; 14: 0.045056; 16: 0.10488; 17: 0.10954
| Rel.Std.Dev. |  | 01: 1.597%; 06: 1.779%; 07: 7.912%; 08: 1.033%; 09: 2.196%; 11: 1.317%; 12: 1.484%; 14: 1.866%; 16: 3.814%; 17: 4.057%
| PDM3 |  | 01: -0.3703%; 06: 0.6565%; 07: 4.841%; 08: 19.97%; 09: -3.72%; 11: 5.213%; 12: -14.95%; 14: -10.11%; 16: 2.359%; 17: 0.4982%
| Z-Score (Absolute) |  | 0.15 | NA | NA | 0.01 | 0.67 | 3.06 | 0.68 | 0.73 | 2.45 | NA | 1.58 | NA | 0.28 | 0.01
| Analytical results for V in OREAS 600 (Certified Value 45.4 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 48.3055 | 44 | 44 | 48 | 45 | 43 | 47.1393 | 47 | 46 | 45.2254 | 45 | 47 | 41 | 50 | 43 | 45.044 | 39 | 46 | 45 | 50 |  | 1
| 1 | 2 | 48.1513 | 44 | 44 | 48 | 44 | 41 | 46.8983 | 48 | 43 | 44.3725 | 46 | 47 | 40 | 49 | 45 | 43.178 | 43 | 44 | 50 | 50 |  | 0
| 1 | 3 | 47.9327 | 45 | 44 | 46 | 44 | 44 | 47.1854 | 50 | 45 | 40.5923 | 47 | 48 | 41 | 52 | 45 | 42.378 | 44 | 42 | 50 | 50 |  | 16
| 1 | 4 | 47.0401 | 46 | 45 | 47 | 45 | 44 | 48.1576 | 48 | 47 | 44.9209 | 45 | 47 | 40 | 52 | 45 | 43.233 | 44 | 44 | 45 | 50 |  | 45.4222
| 1 | 5 | 48.0827 | 44 | 42 | 46 | 45 | 44 | 48.1829 | 49 | 46 | 45.1366 | 45 | 45 | 42 | 52 | 45 | 41.887 | 45 | 44 | 45 | 45 |  | 59
| 1 | 6 | 49.0265 | 45 | 45 | 49 | 44 | 45 | 48.2673 | 48 | 44 | 43.1366 | 46 | 45 | 42 | 49 | 44 | 42.496 | 45 | 44 | 50 | 45
| Mean |  | 48.0898 | 44.6667 | 44 | 47.3333 | 44.5 | 43.5 | 47.6385 | 48.3333 | 45.1667 | 43.8974 | 45.6667 | 46.5 | 41 | 50.6667 | 44.5 | 43.036 | 43.3333 | 44 | 47.5 | 48.3333
| Median |  | 48.117 | 44.5 | 44 | 47.5 | 44.5 | 44 | 47.6715 | 48 | 45.5 | 44.6467 | 45.5 | 47 | 41 | 51 | 45 | 42.837 | 44 | 44 | 47.5 | 50
| Std Dev. |  | 0.641 | 0.8165 | 1.0954 | 1.2111 | 0.54772 | 1.3784 | 0.62668 | 1.0328 | 1.472 | 1.793 | 0.8165 | 1.2247 | 0.89443 | 1.5055 | 0.83666 | 1.1078 | 2.2509 | 1.2649 | 2.7386 | 2.582
| Rel.Std.Dev. |  | 1.333% | 1.828% | 2.49% | 2.559% | 1.231% | 3.169% | 1.315% | 2.137% | 3.259% | 4.085% | 1.788% | 2.634% | 2.182% | 2.971% | 1.88% | 2.574% | 5.194% | 2.875% | 5.766% | 5.342%
| PDM3 |  | 5.873% | -1.663% | -3.131% | 4.207% | -2.03% | -4.232% | 4.879% | 6.409% | -0.5627% | -3.357% | 0.5381% | 2.373% | -9.736% | 11.55% | -2.03% | -5.253% | -4.599% | -3.131% | 4.574% | 6.409%
| Z-Score (Absolute) |  | 1.71 | 0.13 | 0.49 | 1.3 | 0.22 | 0.49 | 1.47 | 1.84 | 0.13 | 0.19 | 0.4 | 0.85 | 2.11 | 3.1 | 0.22 | 1.01 | 0.39 | 0.49 | 1.39 | 1.84
| Analytical results for W in OREAS 600 (Certified Value 4.2 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | < 5 | 3.6 | < 10 | < 10 | < 10 | < 10 | 3.9747 | < 10 | < 0.1 | 2.93 | 5 | < 10 | 13 | 5.4 | < 20 | 5.34 | < 4 | 4.5 | 3.5 |  |  | 1
| 1 | 2 | < 5 | 3.9 | < 10 | 10 | < 10 | < 10 | 3.851 | < 10 | < 0.1 | 3.107 | 4.9 | < 10 | < 10 | 5.4 | < 20 | 5.01 | 4 | 4.5 | 3.5 |  |  | 9
| 1 | 3 | < 5 | 3.9 | < 10 | < 10 | < 10 | < 10 | 3.8264 | < 10 | 2.5 | 3.111 | 4 | < 10 | < 10 | 4.9 | < 20 | 4.37 | < 4 | 4.5 | 4 |  |  | 16
| 1 | 4 | < 5 | 4.1 | < 10 | 10 | < 10 | < 10 | 3.7572 | < 10 | < 0.1 | 2.9834 | 4 | < 10 | < 10 | 4.8 | < 20 | 4.86 | < 4 | 4.5 | 3.5 |  |  | 4.1975
| 1 | 5 | < 5 | 4 | < 10 | 10 | < 10 | < 10 | 3.9228 | < 10 | < 0.1 | 3.0507 | 4.2 | < 10 | < 10 | 4.3 | < 20 | 4.64 | 5 | 4.5 | 3.5 |  |  | 60
| 1 | 6 | < 5 | 4 | < 10 | < 10 | < 10 | < 10 | 3.825 | < 10 | < 0.1 | 3.0257 | 4.9 | < 10 | < 10 | 4.6 | < 20 | 4.8 | 5 | 4.5 | 3.5
| Mean |  |  | 3.9167 |  | 10 |  |  | 3.8595 |  | 2.5 | 3.0346 | 4.5 |  | 13 | 4.9 |  | 4.8367 | 4.6667 | 4.5 | 3.5833
| Median |  |  | 3.95 |  | 10 |  |  | 3.8387 |  | 2.5 | 3.0382 | 4.55 |  | 13 | 4.85 |  | 4.83 | 5 | 4.5 | 3.5
| Std Dev. |  | 02: 0.17224; 06: 0; 09: 0.077558; 13: 0.070662; 14: 0.48166; 17: 0.43818; 19: 0.32916; 23: 0.57735; 24: 0; 25: 0.20412
| Rel.Std.Dev. |  | 02: 4.398%; 06: 0%; 09: 2.01%; 13: 2.329%; 14: 10.7%; 17: 8.942%; 19: 6.806%; 23: 12.37%; 24: 0%; 25: 5.696%
| PDM3 |  |  | -6.69% |  | 138.2% |  |  | -8.052% |  | -40.44% | -27.7% | 7.207% |  | 209.7% | 16.74% |  | 15.23% | 11.18% | 7.207% | -14.63%
| Z-Score (Absolute) |  | NA | 0.6 | NA | 6.39 | NA | NA | 0.74 | NA | 2.32 | 1.7 | 0 | NA | 9.88 | 0.46 | NA | 0.39 | 0.19 | 0 | 1.16
| Analytical results for Y in OREAS 600 (Certified Value 9.93 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | [5 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 10.4722 | 9.46 | 9.6811 | 9 | 10.9 | 12.092 | 9.5 | 10 | < 10 | 10.6 | 8 | 10.03 | 13 | 8 | 9.4 | 9 | [5 empty] | 1
| 1 | 2 | 10.8065 | 9.66 | 9.941 | 10 | 8.4 | 11.02 | 9.8 | 10 | < 10 | 10.8 | 8 | 10.29 | 13 | 8 | 10 | 9.2 | [5 empty] | 0
| 1 | 3 | 10.3401 | 9.99 | 9.7618 | 10 | 10 | 12.6571 | 9.1 | 10 | < 10 | 10.1 | 8 | 10 | 13 | 8 | 9.9 | 9.1 | [5 empty] | 16
| 1 | 4 | 10.2774 | 9.65 | 10.2113 | 9 | 10.7 | 11.2686 | 9.3 | 10 | < 10 | 10.3 | 8 | 10.05 | 12 | 8 | 9.9 | 9.3 | [5 empty] | 9.9267
| 1 | 5 | 10.3492 | 9.69 | 10.1361 | 10 | 10.2 | 11.4224 | 9.4 | 10 | < 10 | 10.6 | 8 | 10.11 | 13 | 8 | 10 | 9.2 | [5 empty] | 61
| 1 | 6 | 11.1625 | 9.69 | 9.9272 | 9 | 10.1 | 11.0952 | 9.9 | 10 | < 10 | 10.8 | 8 | 10.01 | 13 | 8 | 10.3 | 9.4
| Mean |  | 10.568 | 9.69 | 9.9431 | 9.5 | 10.05 | 11.5926 | 9.5 | 10 |  | 10.5333 | 8 | 10.0817 | 12.8333 | 8 | 9.9167 | 9.2
| Median |  | 10.4107 | 9.675 | 9.9341 | 9.5 | 10.15 | 11.3455 | 9.45 | 10 |  | 10.6 | 8 | 10.04 | 13 | 8 | 9.95 | 9.2
| Std Dev. |  | 0.34749 | 0.17053 | 0.20542 | 0.54772 | 0.88261 | 0.6468 | 0.30332 | 0 |  | 0.28048 | 0 | 0.10926 | 0.40825 | 0 | 0.29269 | 0.14142
| Rel.Std.Dev. |  | 3.288% | 1.76% | 2.066% | 5.766% | 8.782% | 5.579% | 3.193% | 0% |  | 2.663% | 0% | 1.084% | 3.181% | 0% | 2.951% | 1.537%
| PDM3 |  | 6.46% | -2.385% | 0.1649% | -4.299% | 1.242% | 16.78% | -4.299% | 0.7383% |  | 6.111% | -19.41% | 1.561% | 29.28% | -19.41% | -0.1012% | -7.321%
| Z-Score (Absolute) |  | 0.77 | 0.39 | 0 | 0.67 | 0.66 | 2.51 | 0.67 | 0.09 | NA | 0.9 | 2.96 | 0.21 | 4.65 | 2.96 | 0.04 | 1.13
| Analytical results for Yb in OREAS 600 (Certified Value 0.91 ppm) | [22 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | [5 empty] | 1
| No | No | 01: 02; 02: 09; 09: 13; 11: 15; 12: 17; 13: 19; 14: 25 | [5 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 09: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*OES/MS; 14: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 01: 0.9; 02: 0.93551; 09: 0.96677; 11: 1.1; 12: 0.9; 13: 0.84; 14: 0.9 | [5 empty] | 1
| 1 | 2 | 01: 0.82; 02: 0.85657; 09: 0.96105; 11: 1.01; 12: 0.9; 13: 0.85; 14: 0.9 | [5 empty] | 0
| 1 | 3 | 01: 0.86; 02: 0.87598; 09: 0.98816; 11: 1.05; 12: 0.9; 13: 0.85; 14: 0.95 | [5 empty] | 16
| 1 | 4 | 01: 0.88; 02: 0.8983; 09: 1.0036; 11: 1.05; 12: 0.9; 13: 0.84; 14: 0.9 | [5 empty] | 0.91257
| 1 | 5 | 01: 0.85; 02: 0.85615; 09: 0.93824; 11: 0.9; 12: 0.9; 13: 0.84; 14: 0.85 | [5 empty] | 62
| 1 | 6 | 01: 0.87; 02: 0.92267; 09: 0.95508; 11: 0.96; 12: 0.9; 13: 0.84; 14: 0.9
| Mean |  | 01: 0.86333; 02: 0.89086; 09: 0.96881; 11: 1.0117; 12: 0.9; 13: 0.84333; 14: 0.9
| Median |  | 01: 0.865; 02: 0.88714; 09: 0.96391; 11: 1.03; 12: 0.9; 13: 0.84; 14: 0.9
| Std Dev. |  | 01: 0.027325; 02: 0.033653; 09: 0.023544; 11: 0.071949; 12: 0; 13: 0.005164; 14: 0.031623
| Rel.Std.Dev. |  | 01: 3.165%; 02: 3.778%; 09: 2.43%; 11: 7.112%; 12: 0%; 13: 0.6123%; 14: 3.514%
| PDM3 |  | 01: -5.396%; 02: -2.379%; 09: 6.163%; 11: 10.86%; 12: -1.378%; 13: -7.587%; 14: -1.378%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.17; 09: 1.27; 11: 2.05; 12: 0; 13: 1.04; 14: 0.18
| Analytical results for Zn in OREAS 600 (Certified Value 615 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 711.3011 | 584 | 629 | 632 | 668 | 622 | 603 | 590 | 625 | 654 | 619.7518 | 902 | 610 | 542 | 564 | 640 | 587.991 | 460 | 620 | 620 | 615 | 1
| 1 | 2 | 700.3087 | 574 | 644 | 665 | 609 | 576 | 592 | 584 | 624 | 631 | 627.2154 | 830 | 589 | 528 | 576 | 657 | 599.597 | 534 | 628 | 625 | 630 | 0
| 1 | 3 | 696.4796 | 575 | 623 | 618 | 617 | 625 | 610 | 576 | 641 | 624 | 618.0664 | 620 | 605 | 534 | 582 | 646 | 610.727 | 519 | 601 | 590 | 620 | 16
| 1 | 4 | 698.5388 | 581 | 649 | 618 | 612 | 636 | 605 | 608 | 630 | 642 | 616.6692 | 617 | 614 | 553 | 580 | 654 | 602.252 | 541 | 613 | 600 | 610 | 614.5297
| 1 | 5 | 707.5386 | 571 | 610 | 603 | 644 | 604 | 607 | 596 | 640 | 626 | 621.033 | 626 | 618 | 532 | 596 | 663 | 595.223 | 564 | 633 | 590 | 610 | 63
| 1 | 6 | 704.9815 | 574 | 634 | 635 | 643 | 636 | 594 | 608 | 631 | 614 | 622.1314 | 836 | 623 | 542 | 559 | 694 | 589.768 | 545 | 609 | 600 | 615
| Mean |  | 703.1914 | 576.5 | 631.5 | 628.5 | 632.1667 | 616.5 | 601.8333 | 593.6667 | 631.8333 | 631.8333 | 620.8112 | 738.5 | 609.8333 | 538.5 | 576.1667 | 659 | 597.593 | 527.1667 | 617.3333 | 604.1667 | 616.6667
| Median |  | 702.6451 | 574.5 | 631.5 | 625 | 630 | 623.5 | 604 | 593 | 630.5 | 628.5 | 620.3924 | 728 | 612 | 538 | 578 | 655.5 | 597.41 | 537.5 | 616.5 | 600 | 615
| Std Dev. |  | 5.7075 | 4.9295 | 14.2092 | 21.2485 | 23.3016 | 23.063 | 7.2503 | 12.9254 | 7.2503 | 14.2045 | 3.7046 | 131.2033 | 11.9569 | 9.0277 | 13.3029 | 18.9737 | 8.453 | 36.0301 | 12.0111 | 14.9722 | 7.5277
| Rel.Std.Dev. |  | 0.8117% | 0.8551% | 2.25% | 3.381% | 3.686% | 3.741% | 1.205% | 2.177% | 1.147% | 2.248% | 0.5967% | 17.77% | 1.961% | 1.676% | 2.309% | 2.879% | 1.415% | 6.835% | 1.946% | 2.478% | 1.221%
| PDM3 |  | 14.43% | -6.188% | 2.762% | 2.273% | 2.87% | 0.3206% | -2.066% | -3.395% | 2.816% | 2.816% | 1.022% | 20.17% | -0.7642% | -12.37% | -6.243% | 7.236% | -2.756% | -14.22% | 0.4562% | -1.686% | 0.3477%
| Z-Score (Absolute) |  | 3.9 | 1.86 | 0.64 | 0.51 | 0.67 | 0.33 | 0.7 | 1.08 | 0.66 | 0.66 | 0.16 | 5.51 | 0.34 | 3.58 | 1.87 | 1.58 | 0.9 | 3.49 | 0 | 0.6 | 0.03
| Analytical results for Zr in OREAS 600 (Certified Value 94 ppm) | [22 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [6 empty] | 0
1
| 1 | 1 | 96.265 | 95 | 97.1788 | 91 | 89.1 | 102.645 | 96.2 | 102 | 100 | 90.6 | 93 | 94.03 | 82 | 88 | 75 | [6 empty] | 1
| 1 | 2 | 95.43 | 96.1 | 97.5872 | 96 | 77 | 103.453 | 96.6 | 104 | 88 | 88.9 | 95 | 93.69 | 97 | 93 | 78 | [6 empty] | 0
| 1 | 3 | 95.78 | 95 | 97.1662 | 94 | 93.7 | 102.051 | 94.4 | 102 | 88 | 92 | 93 | 88.81 | 96 | 92 | 76 | [6 empty] | 16
| 1 | 4 | 96.495 | 97.1 | 98.9281 | 92 | 95.5 | 104.488 | 92.8 | 103 | 86 | 93.8 | 93 | 93.57 | 95 | 93 | 78 | [6 empty] | 93.7526
| 1 | 5 | 95.765 | 94.8 | 98.4912 | 96 | 95 | 109.3376 | 91.7 | 103 | 89 | 91 | 98 | 91.82 | 98 | 92 | 78 | [6 empty] | 64
| 1 | 6 | 97.61 | 96.8 | 96.5085 | 93 | 91.6 | 107.3 | 95.4 | 102 | 89 | 90.2 | 95 | 102.94 | 96 | 93 | 79
| Mean |  | 96.2242 | 95.8 | 97.6433 | 93.6667 | 90.3167 | 104.8791 | 94.5167 | 102.6667 | 90 | 91.0833 | 94.5 | 94.1433 | 94 | 91.8333 | 77.3333
| Median |  | 96.0225 | 95.55 | 97.383 | 93.5 | 92.65 | 103.9705 | 94.9 | 102.5 | 88.5 | 90.8 | 94 | 93.63 | 96 | 92.5 | 78
| Std Dev. |  | 0.77902 | 1.006 | 0.90596 | 2.0656 | 6.9384 | 2.8606 | 1.9416 | 0.8165 | 5.02 | 1.6738 | 1.9748 | 4.728 | 5.9666 | 1.9408 | 1.5055
| Rel.Std.Dev. |  | 0.8096% | 1.05% | 0.9278% | 2.205% | 7.682% | 2.728% | 2.054% | 0.7953% | 5.578% | 1.838% | 2.09% | 5.022% | 6.347% | 2.113% | 1.947%
| PDM3 |  | 2.636% | 2.184% | 4.15% | -0.09169% | -3.665% | 11.87% | 0.815% | 9.508% | -4.003% | -2.847% | 0.7972% | 0.4167% | 0.2639% | -2.047% | -17.51%
| Z-Score (Absolute) |  | 0.55 | 0.41 | 1 | 0.27 | 0.48 | 3.31 | 0.01 | 2.6 | 2.07 | 1.09 | 0 | 0.67 | 0.61 | 0.85 | 5.47
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 600 (Indicative Value 0.488 wt.%) | [17 empty] | Indicative
C | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 02 | 21 | [14 empty] | wt.%
|  |  | IRC | IRC | [14 empty] | 3
3
| 1 | 1 | 0.49 | 0.48597 | [14 empty] | 1
| 1 | 2 | 0.49 | 0.4772 | [14 empty] | 11
| 1 | 3 | 0.48 | 0.49537 | [14 empty] | 16
| 1 | 4 | 0.51 | 0.46946 | [14 empty] | 0.48818
| 1 | 5 | 0.5 | 0.49369 | [14 empty] | 17
| 1 | 6 | 0.49 | 0.47643
| Mean |  | 0.49333 | 0.48302
| Median |  | 0.49 | 0.48159
| Std Dev. |  | 0.010328 | 0.010357
| Rel.Std.Dev. |  | 2.094% | 2.144%
| PDM3 |  | 1.056% | -1.056%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 600 (Certified Value 1.67 wt.%) | [17 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 14 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 1.69 | 1.72 | 1.69 | 1.58 | 1.635 | 1.64 | 1.78 | 1.66 | 1.64 | 1.738 | 1.65 | 1.8854 |  | 1.69 | 1.65 | 1.62 | 1
| 1 | 2 | 1.67 | 1.7 | 1.69 | 1.7 | 1.658 | 1.63 | 1.79 | 1.66 | 1.61 | 1.738 | 1.68 | 1.964 |  | 1.65 | 1.66 | 1.6 | 0
| 1 | 3 | 1.68 | 1.71 | 1.71 | 1.68 | 1.666 | 1.64 | 1.8 | 1.7 | 1.62 | 1.738 | 1.67 | 1.9795 |  | 1.65 | 1.71 | 1.61 | 16
| 1 | 4 | 1.7 | 1.7 | 1.68 | 1.67 | 1.656 | 1.62 | 1.81 | 1.7 | 1.58 | 1.749 | 1.68 | 1.9531 | 1.41 | 1.67 | 1.73 | 1.61 | 1.6686
| 1 | 5 | 1.71 | 1.7 | 1.68 | 1.67 | 1.653 | 1.59 | 1.8 | 1.72 | 1.6 | 1.716 | 1.69 | 2.0011 | 1.47 | 1.65 | 1.66 | 1.6 | 122
| 1 | 6 | 1.71 | 1.72 | 1.67 | 1.68 | 1.657 | 1.64 | 1.81 | 1.66 | 1.61 | 1.749 | 1.71 | 1.9712 | 1.5 | 1.62 | 1.66 | 1.61
| Mean |  | 1.6933 | 1.7083 | 1.6867 | 1.6633 | 1.6542 | 1.6267 | 1.7983 | 1.6833 | 1.61 | 1.738 | 1.68 | 1.959 | 1.46 | 1.655 | 1.6783 | 1.6083
| Median |  | 1.695 | 1.705 | 1.685 | 1.675 | 1.6565 | 1.635 | 1.8 | 1.68 | 1.61 | 1.738 | 1.68 | 1.9676 | 1.47 | 1.65 | 1.66 | 1.61
| Std Dev. |  | 0.01633 | 0.0098319 | 0.013663 | 0.042269 | 0.010342 | 0.019664 | 0.01169 | 0.026583 | 0.02 | 0.01205 | 0.02 | 0.03956 | 0.045826 | 0.023452 | 0.033116 | 0.0075277
| Rel.Std.Dev. |  | 0.9644% | 0.5755% | 0.81% | 2.541% | 0.6252% | 1.209% | 0.6501% | 1.579% | 1.242% | 0.6933% | 1.19% | 2.019% | 3.139% | 1.417% | 1.973% | 0.468%
| PDM3 |  | 1.482% | 2.381% | 1.083% | -0.3158% | -0.8651% | -2.513% | 7.775% | 0.8828% | -3.512% | 4.159% | 0.6831% | 17.41% | -12.5% | -0.8152% | 0.5832% | -3.612%
| Z-Score (Absolute) |  | 0.33 | 0.71 | 0.17 | 0 | 0.64 | 1.33 | 2.95 | 0.08 | 1.74 | 1.44 | 0 | 7.31 | 5.48 | 0.62 | 0.3 | 1.78
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.