## Abbreviations
Table 1. Abbreviations used for OREAS 604
Abbreviation | Explanation
Std.Dev. | one sigma standard deviation
Rel.Std.Dev. | one sigma relative standard deviation
PDM3 | percent deviation of lab mean from corrected mean of means
NR | not reported
< | ‘less than’
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with AAS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with either AAS or ICP-OES finish as appropriate
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/AAS | ore grade aqua regia digest with either AAS or ICP-OES finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
FA*AAS | fire assay with AAS finish
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
NB. Formatted area for below table ends at row 46 inclusive, allowing for 40 maximum potential laboratories.
Table 5. Participating Laboratory List used for OREAS 604
Lab Slot | Laboratory Name and Location
1.21 | Accurassay, Thunder Bay, Ontario, Canada
1.23 | Acme (BV), Santiago, Chile
2.01 | Actlabs, Ancaster, Ontario, Canada
1.2 | AH Knight, Spartanburg, SC, USA
1.06 | ALS, Johannesburg, South Africa
1.05 | ALS, Lima, Peru
1.03 | ALS, Reno, Nevada, USA
1.08 | ALS, Townsville, QLD, Australia
1.04 | ALS, Val-d'or, Quebec, Canada
1.07 | ALS, Vancouver, BC, Canada
1.24 | Bureau Veritas Geoanalytical, Adelaide, SA, Australia
1.27 | Bureau Veritas Geoanalytical, Kalgoorlie, WA, Australia
1.25 | Bureau Veritas Geoanalytical, Perth, WA, Australia
1.26 | Bureau Veritas Kalassay, Kalgoorlie, WA, Australia
1.11 | Inspectorate (BV), Lima, Peru
1.12 | Inspectorate (BV), Sparks, Nevada, USA
1.02 | Intertek Genalysis, Adelaide, SA, Australia
1.14 | Intertek Genalysis, Perth, WA, Australia
1.09 | Intertek Testing Services, Cupang, Muntinlupa, Philippines
1.15 | Intertek Testing Services, Shunyi, Beijing, China
1.22 | MINTEK Analytical Services, Randburg, South Africa
1.01 | PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
1.16 | SGS de Mexico, Durango, Mexico
1.18 | SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
1.19 | SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
1.17 | SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
1.13 | Shiva Analyticals Ltd, Bangalore North, Karnataka, India
1.1 | SRL, Perth, WA, Australia

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 604
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Ag, Silver (ppm) | 488.4652 | 14.5635 | 480.6537 | 496.2766 | 479.5997 | 497.3306
Au, Gold (ppm) | 1.427 | 0.055195 | 1.4072 | 1.4469 | 1.4103 | 1.4438
4-Acid Digestion
Ag, Silver (ppm) | 490.6509 | 10.2864 | 486.0278 | 495.2741 | 482.586 | 498.7158
Al, Aluminium (wt.%) | 5.8168 | 0.23555 | 5.6948 | 5.9387 | 5.6916 | 5.9419
As, Arsenic (ppm) | 971.931 | 62.0951 | 943.0295 | 1000.8326 | 945.1586 | 998.7035
Be, Beryllium (ppm) | 1.125 | 0.11093 | 1.0748 | 1.1752 | 1.0591 | 1.1909
Bi, Bismuth (ppm) | 28.3856 | 4.1895 | 26.2759 | 30.4952 | 27.1737 | 29.5975
Ca, Calcium (wt.%) | 0.73504 | 0.024102 | 0.72355 | 0.74653 | 0.71233 | 0.75774
Cd, Cadmium (ppm) | 14.8721 | 1.063 | 14.3617 | 15.3826 | 14.2971 | 15.4472
Ce, Cerium (ppm) | 38.0843 | 4.0298 | 34.639 | 41.5295 | 35.7616 | 40.4069
Co, Cobalt (ppm) | 41.971 | 2.8419 | 40.6916 | 43.2504 | 40.8408 | 43.1012
Cr, Chromium (ppm) | 33.8189 | 3.5765 | 32.2842 | 35.3535 | 30.5802 | 37.0575
Cs, Cesium (ppm) | 3.6346 | 0.18723 | 3.5082 | 3.7609 | 3.4709 | 3.7983
Cu, Copper (wt.%) | 2.1641 | 0.048518 | 2.1451 | 2.1831 | 2.1274 | 2.2008
Dy, Dysprosium (ppm) | 1.6008 | 0.10024 | 1.4999 | 1.7018 | 1.5111 | 1.6906
Er, Erbium (ppm) | 0.62428 | 0.043727 | 0.58754 | 0.66102 | 0.5618 | 0.68676
Eu, Europium (ppm) | 0.73786 | 0.11288 | 0.59014 | 0.88558 | 0.6754 | 0.80031
Fe, Iron (wt.%) | 3.0216 | 0.2099 | 2.9272 | 3.116 | 2.9681 | 3.0751
Ga, Gallium (ppm) | 26.7033 | 2.6264 | 25.3877 | 28.0189 | 26.087 | 27.3196
Gd, Gadolinium (ppm) | 2.7195 | 0.29562 | 2.3258 | 3.1132 | 2.4889 | 2.9501
Ge, Germanium (ppm) | < 5 | IND | IND | IND | IND | IND
Hf, Hafnium (ppm) | 3.148 | 0.22895 | 3.0039 | 3.2921 | 3.0142 | 3.2819
Ho, Holmium (ppm) | 0.2461 | 0.028242 | 0.21658 | 0.27563 | 0.22891 | 0.2633
In, Indium (ppm) | 3.5749 | 0.29085 | 3.3688 | 3.781 | 3.4976 | 3.6522
K, Potassium (wt.%) | 1.3181 | 0.058557 | 1.2934 | 1.3429 | 1.2932 | 1.3431
La, Lanthanum (ppm) | 19.3687 | 3.0747 | 16.947 | 21.7904 | 18.3779 | 20.3595
Li, Lithium (ppm) | 22.6336 | 1.8532 | 21.4984 | 23.7688 | 21.9286 | 23.3386
Lu, Lutetium (ppb) | 80.7266 | 13.135 | 73.868 | 87.5851 | IND | IND
Mg, Magnesium (wt.%) | 0.20768 | 0.01293 | 0.20168 | 0.21369 | 0.20166 | 0.21371
Mn, Manganese (ppm) | 243.748 | 12.8517 | 237.9708 | 249.5253 | 237.8132 | 249.6829
Mo, Molybdenum (ppm) | 4.1145 | 0.57978 | 3.8649 | 4.3642 | 3.8902 | 4.3389
Na, Sodium (wt.%) | 0.83567 | 0.033623 | 0.81969 | 0.85165 | 0.82106 | 0.85028
Nb, Niobium (ppm) | 8.8249 | 0.70026 | 8.3284 | 9.3215 | 8.4276 | 9.2223
Nd, Neodymium (ppm) | 16.5097 | 1.0789 | 15.3807 | 17.6388 | 15.6266 | 17.3929
Ni, Nickel (ppm) | 637.5076 | 30.5042 | 623.9403 | 651.0749 | 623.92 | 651.0952
P, Phosphorus (ppm) | 572.5113 | 46.8406 | 548.4972 | 596.5254 | 549.8411 | 595.1815
Pb, Lead (ppm) | 994.348 | 69.063 | 961.6479 | 1027.0482 | 972.8035 | 1015.8926
Pr, Praseodymium (ppm) | 4.5488 | 0.35847 | 4.1628 | 4.9348 | 4.2781 | 4.8196
Rb, Rubidium (ppm) | 51.0422 | 2.7857 | 48.9246 | 53.1597 | 49.7841 | 52.3003
S, Sulphur (wt.%) | 4.592 | 0.20113 | 4.4918 | 4.6921 | 4.5095 | 4.6745
Sb, Antimony (ppm) | 166.5979 | 7.6524 | 162.8388 | 170.3571 | 161.7621 | 171.4338
Sc, Scandium (ppm) | 4.8145 | 0.64207 | 4.5138 | 5.1151 | 4.54 | 5.0889
Se, Selenium (ppm) | 43.8515 | 5.5377 | 40.1489 | 47.5542 | 41.0121 | 46.691
Sm, Samarium (ppm) | 3.0831 | 0.1904 | 2.8843 | 3.282 | 2.9861 | 3.1802
Sn, Tin (ppm) | 3.8331 | 0.26068 | 3.6178 | 4.0484 | 3.6236 | 4.0426
Sr, Strontium (ppm) | 397.8423 | 37.5301 | 378.8007 | 416.8838 | 382.2633 | 413.4213
Ta, Tantalum (ppm) | < 1 | IND | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.32573 | 0.029847 | 0.29742 | 0.35404 | 0.29686 | 0.3546
Te, Tellurium (ppm) | 25.1077 | 3.7474 | 21.5922 | 28.6232 | 23.5415 | 26.6739
Th, Thorium (ppm) | 7.4708 | 0.9024 | 6.8005 | 8.1412 | 6.8769 | 8.0648
Ti, Titanium (wt.%) | 0.19057 | 0.01473 | 0.18338 | 0.19777 | 0.18523 | 0.19592
Tl, Thallium (ppm) | 7.4435 | 0.82601 | 6.947 | 7.9401 | 7.187 | 7.7001
Tm, Thulium (ppb) | 81.1841 | 13.6685 | 73.2833 | 89.0848 | IND | IND
U, Uranium (ppm) | 3.1653 | 0.14205 | 3.0724 | 3.2582 | 3.0261 | 3.3045
V, Vanadium (ppm) | 36.0056 | 1.8938 | 35.1149 | 36.8963 | 34.7978 | 37.2133
W, Tungsten (ppm) | 16.7237 | 2.8511 | 14.8982 | 18.5493 | 15.9644 | 17.4831
Y, Yttrium (ppm) | 7.1644 | 0.29497 | 6.9953 | 7.3335 | 6.8799 | 7.4489
Yb, Ytterbium (ppm) | 0.5631 | 0.036583 | 0.53445 | 0.59175 | 0.51789 | 0.60831
Zn, Zinc (wt.%) | 0.25485 | 0.0082314 | 0.25118 | 0.25851 | 0.24992 | 0.25978
Zr, Zirconium (ppm) | 104.4469 | 7.038 | 100.6839 | 108.21 | 101.4371 | 107.4568
Aqua Regia Digestion
Ag, Silver (ppm) | 492.363 | 15.2289 | 484.0967 | 500.6294 | 481.9042 | 502.8219
Al, Aluminium (wt.%) | 0.78979 | 0.1345 | 0.72722 | 0.85236 | 0.76556 | 0.81402
As, Arsenic (ppm) | 966.0952 | 52.5012 | 941.1399 | 991.0505 | 946.8399 | 985.3504
Au, Gold (ppm) | 1.428 | 0.056053 | 1.3951 | 1.4609 | 1.4096 | 1.4464
B, Boron (ppm) | < 10 | IND | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.33436 | 0.057821 | 0.27788 | 0.39084 | 0.30234 | 0.36638
Bi, Bismuth (ppm) | 26.9191 | 4.7087 | 24.6058 | 29.2323 | 25.7081 | 28.1301
Ca, Calcium (wt.%) | 0.59112 | 0.037189 | 0.57445 | 0.6078 | 0.57798 | 0.60427
Cd, Cadmium (ppm) | 15.5259 | 1.3539 | 14.8903 | 16.1615 | 15.0671 | 15.9846
Ce, Cerium (ppm) | 23.5183 | 3.2935 | 20.8802 | 26.1564 | 22.4874 | 24.5491
Co, Cobalt (ppm) | 42.5537 | 2.4055 | 41.4976 | 43.6098 | 41.1359 | 43.9714
Cr, Chromium (ppm) | 32.0819 | 2.9198 | 31.0076 | 33.1562 | 29.9604 | 34.2034
Cs, Cesium (ppm) | 1.2226 | 0.22274 | 1.0313 | 1.414 | 1.136 | 1.3092
Cu, Copper (wt.%) | 2.1579 | 0.063901 | 2.1305 | 2.1854 | 2.1146 | 2.2013
Dy, Dysprosium (ppm) | 0.86204 | 0.10744 | 0.74377 | 0.98031 | 0.82839 | 0.89569
Er, Erbium (ppm) | 0.28977 | 0.032878 | 0.24738 | 0.33217 | 0.27347 | 0.30608
Eu, Europium (ppm) | 0.38892 | 0.057776 | 0.31213 | 0.46572 | 0.36574 | 0.41211
Fe, Iron (wt.%) | 2.9267 | 0.14887 | 2.8613 | 2.9922 | 2.8592 | 2.9943
Ga, Gallium (ppm) | 6.0957 | 0.84565 | 5.5549 | 6.6365 | 5.9007 | 6.2907
Gd, Gadolinium (ppm) | 1.4851 | 0.20844 | 1.2521 | 1.7181 | 1.4083 | 1.5618
Hf, Hafnium (ppm) | 0.57778 | 0.11451 | 0.48579 | 0.66978 | 0.53784 | 0.61773
Hg, Mercury (ppm) | 1.0073 | 0.051829 | 0.9764 | 1.0382 | 0.92097 | 1.0937
Ho, Holmium (ppm) | 0.1393 | 0.0063238 | 0.13253 | 0.14608 | IND | IND
In, Indium (ppm) | 3.6626 | 0.40487 | 3.3448 | 3.9804 | 3.5349 | 3.7902
K, Potassium (wt.%) | 0.16692 | 0.020968 | 0.15622 | 0.17762 | 0.15592 | 0.17791
La, Lanthanum (ppm) | 10.5317 | 1.8664 | 9.5597 | 11.5038 | 9.885 | 11.1784
Li, Lithium (ppm) | 6.1564 | 0.84304 | 5.6213 | 6.6915 | 5.8828 | 6.43
Lu, Lutetium (ppb) | 27.3341 | 4.2858 | 23.271 | 31.3973 | IND | IND
Mg, Magnesium (wt.%) | 0.10638 | 0.014388 | 0.099864 | 0.11289 | IND | IND
Mn, Manganese (ppm) | 230.9176 | 8.4502 | 226.914 | 234.9213 | 224.8113 | 237.024
Mo, Molybdenum (ppm) | 4.0186 | 0.36078 | 3.8678 | 4.1694 | 3.7761 | 4.261
Na, Sodium (ppm) | 461.4667 | 84.0248 | 422.1276 | 500.8057 | 450.1092 | 472.8241
Nb, Niobium (ppm) | 0.38918 | 0.075254 | 0.28897 | 0.4894 | 0.34372 | 0.43465
Nd, Neodymium (ppm) | 10.5923 | 1.4128 | 9.0078 | 12.1769 | 10.2543 | 10.9304
Ni, Nickel (ppm) | 649.9844 | 27.6666 | 637.0378 | 662.931 | 635.1764 | 664.7925
P, Phosphorus (ppm) | 247.7249 | 31.0927 | 230.6678 | 264.7821 | 235.8575 | 259.5924
Pb, Lead (ppm) | 703.095 | 36.5767 | 687.1389 | 719.0511 | 687.8316 | 718.3584
Pr, Praseodymium (ppm) | 2.8257 | 0.42958 | 2.3441 | 3.3073 | 2.7076 | 2.9437
Rb, Rubidium (ppm) | 8.3917 | 1.3912 | 7.0344 | 9.7491 | 7.9314 | 8.852
S, Sulphur (wt.%) | 4.258 | 0.21269 | 4.1531 | 4.3628 | 4.1545 | 4.3614
Sb, Antimony (ppm) | 126.2324 | 21.9906 | 115.0966 | 137.3682 | 121.9992 | 130.4656
Sc, Scandium (ppm) | < 2 | IND | IND | IND | IND | IND
Se, Selenium (ppm) | 45.0449 | 5.0859 | 41.9576 | 48.1322 | 43.1825 | 46.9073
Sm, Samarium (ppm) | 1.9443 | 0.2494 | 1.67 | 2.2187 | 1.8494 | 2.0393
Sn, Tin (ppm) | 2.8083 | 0.25115 | 2.5913 | 3.0253 | 2.6409 | 2.9757
Sr, Strontium (ppm) | 37.7147 | 6.0293 | 34.5676 | 40.8619 | 36.3851 | 39.0444
Tb, Terbium (ppm) | 0.18568 | 0.032584 | 0.15781 | 0.21355 | 0.17417 | 0.19719
Te, Tellurium (ppm) | 25.3294 | 1.84 | 23.9182 | 26.7406 | 24.2837 | 26.3751
Th, Thorium (ppm) | 3.7373 | 0.70705 | 3.1899 | 4.2847 | 3.5324 | 3.9422
Ti, Titanium (ppm) | 97.0729 | 19.7967 | 83.7858 | 110.3599 | IND | IND
Tl, Thallium (ppm) | 7.2658 | 1.1582 | 6.5771 | 7.9546 | 7.0085 | 7.5232
U, Uranium (ppm) | 1.215 | 0.16891 | 1.0826 | 1.3474 | 1.1458 | 1.2842
V, Vanadium (ppm) | 10.1531 | 1.5518 | 9.4105 | 10.8957 | 9.6505 | 10.6557
W, Tungsten (ppm) | 3.497 | 0.70911 | 2.941 | 4.053 | 3.2708 | 3.7232
Y, Yttrium (ppm) | 3.447 | 0.50558 | 3.0968 | 3.7973 | 3.3031 | 3.591
Yb, Ytterbium (ppm) | 0.19529 | 0.018443 | 0.18384 | 0.20674 | 0.18205 | 0.20853
Zn, Zinc (wt.%) | 0.2536 | 0.011648 | 0.24791 | 0.2593 | 0.24976 | 0.25744
Zr, Zirconium (ppm) | 21.1335 | 2.9402 | 19.3286 | 22.9385 | 20.1168 | 22.1503
Infrared Combustion
S, Sulphur (wt.%) | 4.8539 | 0.21305 | 4.734 | 4.9738 | 4.7845 | 4.9233

## Indicative Values
Table 3. Indicative Values for OREAS 604
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppm | 5.1917 | Pt | ppb | 895.4167
4-Acid Digestion
B | ppm | 13.3333 | Hg | ppm | 0.66667
Ba | ppm | 5923.2928 | Re | ppb | 2.0634
Aqua Regia Digestion
Ba | ppm | 409.5 | Pt | ppb | 890.2983 | Ta | ppm | < 0.01
Ge | ppm | 0.2035 | Re | ppb | 1.25 | Tm | ppb | 31.1221
Pd | ppm | 5.0124 | Ru | ppb | < 2
Infrared Combustion
C | wt.% | 0.1578

## Indicative Values (2)
Table 3. Indicative Values for OREAS 604
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Laser Ablation ICP-MS
Ag | ppm | 477.5 | Ho | ppm | 0.295 | Sn | ppm | 4.5
As | ppm | 951.5 | In | ppm | 3.45 | Sr | ppm | 480
Ba | ppm | 8380 | La | ppm | 24.95 | Ta | ppm | 0.825
Be | ppm | 0.25 | Lu | ppm | 0.08 | Tb | ppm | 0.35
Bi | ppm | 27.35 | Mn | wt.% | 0.02475 | Te | ppm | 25.4
Cd | ppm | 15.55 | Mo | ppm | 5.1 | Th | ppm | 8.685
Ce | ppm | 43.6 | Nb | ppm | 9.335 | Ti | wt.% | 0.1915
Co | ppm | 44.9 | Nd | ppm | 18.2 | Tl | ppm | 7.6
Cr | ppm | 37.5 | Ni | ppm | 612 | Tm | ppm | 0.11
Cs | ppm | 3.63 | Pb | ppm | 1010.5 | U | ppm | 3.11
Cu | ppm | 20850 | Pr | ppm | 5.04 | V | ppm | 36.35
Dy | ppm | 1.64 | Rb | ppm | 53.3 | W | ppm | 16.2
Er | ppm | 0.635 | Re | ppm | 0.0075 | Y | ppm | 8.18
Eu | ppm | 0.56 | Sb | ppm | 155 | Yb | ppm | 0.62
Ga | ppm | 27.8 | Sc | ppm | 5 | Zn | ppm | 2485
Gd | ppm | 2.645 | Se | ppm | < 5 | Zr | ppm | 146
Hf | ppm | 4.335 | Sm | ppm | 3.72
Borate Fusion XRF
Al2O3 | wt.% | 11.37 | Fe2O3 | wt.% | 4.425 | Pb | ppm | 1045
As | ppm | 960 | K2O | wt.% | 1.625 | SiO2 | wt.% | 66.72
Ba | ppm | 8635 | MgO | wt.% | 0.39 | Sn | ppm | 12.5
CaO | wt.% | 1.025 | MnO | wt.% | 0.04 | SO3 | wt.% | 11.45
Co | ppm | 50 | Na2O | wt.% | 1.135 | TiO2 | wt.% | 0.3275
Cr | ppm | 30 | Ni | ppm | 625 | U | ppm | < 10
Cu | ppm | 21850 | P2O5 | wt.% | 0.1365 | Zn | ppm | 2435
Thermogravimetry
LOI1000 | wt.% | 7.295

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 604
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Ag, ppm | 488.4652 | 14.5635 | 459.3381 | 517.5922 | 444.7746 | 532.1557 | 2.981% | 5.963% | 8.944% | 464.0419 | 512.8884
Au, ppm | 1.427 | 0.055195 | 1.3166 | 1.5374 | 1.2614 | 1.5926 | 3.868% | 7.736% | 11.6% | 1.3557 | 1.4984
4-Acid Digestion
Ag, ppm | 490.6509 | 10.2864 | 470.078 | 511.2238 | 459.7916 | 521.5103 | 2.096% | 4.193% | 6.289% | 466.1184 | 515.1835
Al, wt.% | 5.8168 | 0.23555 | 5.3457 | 6.2879 | 5.1101 | 6.5234 | 4.05% | 8.099% | 12.15% | 5.5259 | 6.1076
As, ppm | 971.931 | 62.0951 | 847.7409 | 1096.1211 | 785.6459 | 1158.2162 | 6.389% | 12.78% | 19.17% | 923.3345 | 1020.5276
Be, ppm | 1.125 | 0.11093 | 0.90317 | 1.3469 | 0.79224 | 1.4578 | 9.86% | 19.72% | 29.58% | 1.0688 | 1.1813
Bi, ppm | 28.3856 | 4.1895 | 20.0066 | 36.7646 | 15.8171 | 40.9541 | 14.76% | 29.52% | 44.28% | 26.9663 | 29.8049
Ca, wt.% | 0.73504 | 0.024102 | 0.68683 | 0.78324 | 0.66273 | 0.80735 | 3.279% | 6.558% | 9.837% | 0.69829 | 0.77179
Cd, ppm | 14.8721 | 1.063 | 12.746 | 16.9982 | 11.683 | 18.0613 | 7.148% | 14.3% | 21.44% | 14.1285 | 15.6157
Ce, ppm | 38.0843 | 4.0298 | 30.0246 | 46.1439 | 25.9947 | 50.1738 | 10.58% | 21.16% | 31.74% | 36.18 | 39.9885
Co, ppm | 41.971 | 2.8419 | 36.2872 | 47.6548 | 33.4453 | 50.4967 | 6.771% | 13.54% | 20.31% | 39.8724 | 44.0695
Cr, ppm | 33.8189 | 3.5765 | 26.6659 | 40.9718 | 23.0894 | 44.5483 | 10.58% | 21.15% | 31.73% | 32.1279 | 35.5098
Cs, ppm | 3.6346 | 0.18723 | 3.2601 | 4.009 | 3.0729 | 4.1963 | 5.151% | 10.3% | 15.45% | 3.4528 | 3.8163
Cu, wt.% | 2.1641 | 0.048518 | 2.0671 | 2.2611 | 2.0186 | 2.3097 | 2.242% | 4.484% | 6.726% | 2.0559 | 2.2723
Dy, ppm | 1.6008 | 0.10024 | 1.4004 | 1.8013 | 1.3001 | 1.9015 | 6.262% | 12.52% | 18.78% | 1.5208 | 1.6809
Er, ppm | 0.62428 | 0.043727 | 0.53683 | 0.71173 | 0.4931 | 0.75546 | 7.004% | 14.01% | 21.01% | 0.59307 | 0.65549
Eu, ppm | 0.73786 | 0.11288 | 0.5121 | 0.96362 | 0.39922 | 1.0765 | 15.3% | 30.6% | 45.89% | 0.70096 | 0.77475
Fe, wt.% | 3.0216 | 0.2099 | 2.6018 | 3.4414 | 2.3919 | 3.6513 | 6.947% | 13.89% | 20.84% | 2.8705 | 3.1727
Ga, ppm | 26.7033 | 2.6264 | 21.4504 | 31.9561 | 18.824 | 34.5825 | 9.836% | 19.67% | 29.51% | 25.3681 | 28.0384
Gd, ppm | 2.7195 | 0.29562 | 2.1283 | 3.3107 | 1.8326 | 3.6064 | 10.87% | 21.74% | 32.61% | 2.5835 | 2.8555
Ge, ppm | < 5 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Hf, ppm | 3.148 | 0.22895 | 2.6901 | 3.6059 | 2.4612 | 3.8349 | 7.273% | 14.55% | 21.82% | 2.9906 | 3.3054
Ho, ppm | 0.2461 | 0.028242 | 0.18962 | 0.30259 | 0.16138 | 0.33083 | 11.48% | 22.95% | 34.43% | 0.2338 | 0.25841
In, ppm | 3.5749 | 0.29085 | 2.9932 | 4.1566 | 2.7023 | 4.4474 | 8.136% | 16.27% | 24.41% | 3.3961 | 3.7536
K, wt.% | 1.3181 | 0.058557 | 1.201 | 1.4353 | 1.1425 | 1.4938 | 4.442% | 8.885% | 13.33% | 1.2522 | 1.384
La, ppm | 19.3687 | 3.0747 | 13.2193 | 25.5181 | 10.1446 | 28.5928 | 15.87% | 31.75% | 47.62% | 18.4003 | 20.3371
Li, ppm | 22.6336 | 1.8532 | 18.9273 | 26.3399 | 17.0741 | 28.1931 | 8.188% | 16.38% | 24.56% | 21.5019 | 23.7653
Lu, ppb | 80.7266 | 13.135 | 54.4565 | 106.9966 | 41.3214 | 120.1317 | 16.27% | 32.54% | 48.81% | 76.6902 | 84.7629
Mg, wt.% | 0.20768 | 0.01293 | 0.18182 | 0.23354 | 0.16889 | 0.24647 | 6.226% | 12.45% | 18.68% | 0.1973 | 0.21807
Mn, ppm | 243.748 | 12.8517 | 218.0446 | 269.4514 | 205.1929 | 282.3031 | 5.273% | 10.55% | 15.82% | 231.5606 | 255.9354
Mo, ppm | 4.1145 | 0.57978 | 2.955 | 5.2741 | 2.3752 | 5.8539 | 14.09% | 28.18% | 42.27% | 3.9088 | 4.3203
Na, wt.% | 0.83567 | 0.033623 | 0.76843 | 0.90292 | 0.7348 | 0.93654 | 4.023% | 8.047% | 12.07% | 0.79389 | 0.87746
Nb, ppm | 8.8249 | 0.70026 | 7.4244 | 10.2255 | 6.7242 | 10.9257 | 7.935% | 15.87% | 23.8% | 8.3837 | 9.2662
Nd, ppm | 16.5097 | 1.0789 | 14.352 | 18.6674 | 13.2732 | 19.7463 | 6.535% | 13.07% | 19.6% | 15.6842 | 17.3352
Ni, ppm | 637.5076 | 30.5042 | 576.4991 | 698.5161 | 545.9948 | 729.0203 | 4.785% | 9.57% | 14.35% | 605.6322 | 669.383
P, ppm | 572.5113 | 46.8406 | 478.8302 | 666.1924 | 431.9897 | 713.033 | 8.182% | 16.36% | 24.54% | 543.8857 | 601.1369
Pb, ppm | 994.348 | 69.063 | 856.2221 | 1132.474 | 787.1592 | 1201.5369 | 6.946% | 13.89% | 20.84% | 944.6306 | 1044.0654
Pr, ppm | 4.5488 | 0.35847 | 3.8319 | 5.2658 | 3.4734 | 5.6242 | 7.881% | 15.76% | 23.64% | 4.3214 | 4.7763
Rb, ppm | 51.0422 | 2.7857 | 45.4708 | 56.6136 | 42.685 | 59.3993 | 5.458% | 10.92% | 16.37% | 48.4901 | 53.5943
S, wt.% | 4.592 | 0.20113 | 4.1897 | 4.9942 | 3.9886 | 5.1954 | 4.38% | 8.76% | 13.14% | 4.3624 | 4.8216
Sb, ppm | 166.5979 | 7.6524 | 151.2931 | 181.9027 | 143.6407 | 189.5551 | 4.593% | 9.187% | 13.78% | 158.268 | 174.9278
Sc, ppm | 4.8145 | 0.64207 | 3.5303 | 6.0986 | 2.8882 | 6.7407 | 13.34% | 26.67% | 40.01% | 4.5737 | 5.0552
Se, ppm | 43.8515 | 5.5377 | 32.7762 | 54.9268 | 27.2386 | 60.4645 | 12.63% | 25.26% | 37.88% | 41.659 | 46.0441
Sm, ppm | 3.0831 | 0.1904 | 2.7023 | 3.4639 | 2.5119 | 3.6543 | 6.175% | 12.35% | 18.53% | 2.929 | 3.2373
Sn, ppm | 3.8331 | 0.26068 | 3.3117 | 4.3544 | 3.0511 | 4.6151 | 6.801% | 13.6% | 20.4% | 3.6414 | 4.0247
Sr, ppm | 397.8423 | 37.5301 | 322.782 | 472.9025 | 285.2519 | 510.4326 | 9.433% | 18.87% | 28.3% | 377.9502 | 417.7344
Ta, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.32573 | 0.029847 | 0.26604 | 0.38542 | 0.23619 | 0.41527 | 9.163% | 18.33% | 27.49% | 0.30944 | 0.34202
Te, ppm | 25.1077 | 3.7474 | 17.6129 | 32.6025 | 13.8655 | 36.3499 | 14.93% | 29.85% | 44.78% | 23.8523 | 26.3631
Th, ppm | 7.4708 | 0.9024 | 5.666 | 9.2756 | 4.7636 | 10.178 | 12.08% | 24.16% | 36.24% | 7.0973 | 7.8444
Ti, wt.% | 0.19057 | 0.01473 | 0.16111 | 0.22003 | 0.14638 | 0.23476 | 7.729% | 15.46% | 23.19% | 0.18105 | 0.2001
Tl, ppm | 7.4435 | 0.82601 | 5.7915 | 9.0956 | 4.9655 | 9.9216 | 11.1% | 22.19% | 33.29% | 7.0714 | 7.8157
Tm, ppb | 81.1841 | 13.6685 | 53.8471 | 108.521 | 40.1786 | 122.1895 | 16.84% | 33.67% | 50.51% | 77.1249 | 85.2433
U, ppm | 3.1653 | 0.14205 | 2.8812 | 3.4494 | 2.7391 | 3.5914 | 4.488% | 8.975% | 13.46% | 3.007 | 3.3236
V, ppm | 36.0056 | 1.8938 | 32.218 | 39.7932 | 30.3241 | 41.687 | 5.26% | 10.52% | 15.78% | 34.2053 | 37.8058
W, ppm | 16.7237 | 2.8511 | 11.0215 | 22.426 | 8.1703 | 25.2772 | 17.05% | 34.1% | 51.15% | 15.8875 | 17.5599
Y, ppm | 7.1644 | 0.29497 | 6.5744 | 7.7543 | 6.2795 | 8.0493 | 4.117% | 8.234% | 12.35% | 6.8062 | 7.5226
Yb, ppm | 0.5631 | 0.036583 | 0.48993 | 0.63627 | 0.45335 | 0.67285 | 6.497% | 12.99% | 19.49% | 0.53495 | 0.59126
Zn, wt.% | 0.25485 | 0.0082314 | 0.23838 | 0.27131 | 0.23015 | 0.27954 | 3.23% | 6.46% | 9.69% | 0.2421 | 0.26759
Zr, ppm | 104.4469 | 7.038 | 90.3709 | 118.5229 | 83.3329 | 125.5609 | 6.738% | 13.48% | 20.22% | 99.2246 | 109.6693
Aqua Regia Digestion
Ag, ppm | 492.363 | 15.2289 | 461.9053 | 522.8208 | 446.6765 | 538.0496 | 3.093% | 6.186% | 9.279% | 467.7449 | 516.9812
Al, wt.% | 0.78979 | 0.1345 | 0.52078 | 1.0588 | 0.38627 | 1.1933 | 17.03% | 34.06% | 51.09% | 0.7503 | 0.82928
As, ppm | 966.0952 | 52.5012 | 861.0927 | 1071.0976 | 808.5915 | 1123.5989 | 5.434% | 10.87% | 16.3% | 917.7904 | 1014.3999
Au, ppm | 1.428 | 0.056053 | 1.3159 | 1.5401 | 1.2598 | 1.5962 | 3.925% | 7.851% | 11.78% | 1.3566 | 1.4994
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.33436 | 0.057821 | 0.21872 | 0.45 | 0.1609 | 0.50782 | 17.29% | 34.59% | 51.88% | 0.31764 | 0.35108
Bi, ppm | 26.9191 | 4.7087 | 17.5017 | 36.3365 | 12.793 | 41.0452 | 17.49% | 34.98% | 52.48% | 25.5731 | 28.265
Ca, wt.% | 0.59112 | 0.037189 | 0.51675 | 0.6655 | 0.47956 | 0.70269 | 6.291% | 12.58% | 18.87% | 0.56157 | 0.62068
Cd, ppm | 15.5259 | 1.3539 | 12.8181 | 18.2337 | 11.4642 | 19.5876 | 8.72% | 17.44% | 26.16% | 14.7496 | 16.3022
Ce, ppm | 23.5183 | 3.2935 | 16.9313 | 30.1052 | 13.6378 | 33.3987 | 14% | 28.01% | 42.01% | 22.3424 | 24.6942
Co, ppm | 42.5537 | 2.4055 | 37.7426 | 47.3647 | 35.3371 | 49.7703 | 5.653% | 11.31% | 16.96% | 40.426 | 44.6814
Cr, ppm | 32.0819 | 2.9198 | 26.2423 | 37.9215 | 23.3224 | 40.8413 | 9.101% | 18.2% | 27.3% | 30.4778 | 33.686
Cs, ppm | 1.2226 | 0.22274 | 0.77716 | 1.6681 | 0.55442 | 1.8909 | 18.22% | 36.44% | 54.65% | 1.1615 | 1.2838
Cu, wt.% | 2.1579 | 0.063901 | 2.0301 | 2.2857 | 1.9662 | 2.3496 | 2.961% | 5.922% | 8.884% | 2.05 | 2.2658
Dy, ppm | 0.86204 | 0.10744 | 0.64717 | 1.0769 | 0.53973 | 1.1844 | 12.46% | 24.93% | 37.39% | 0.81894 | 0.90514
Er, ppm | 0.28977 | 0.032878 | 0.22402 | 0.35553 | 0.19114 | 0.38841 | 11.35% | 22.69% | 34.04% | 0.27528 | 0.30426
Eu, ppm | 0.38892 | 0.057776 | 0.27337 | 0.50448 | 0.2156 | 0.56225 | 14.86% | 29.71% | 44.57% | 0.36948 | 0.40837
Fe, wt.% | 2.9267 | 0.14887 | 2.629 | 3.2245 | 2.4801 | 3.3734 | 5.087% | 10.17% | 15.26% | 2.7804 | 3.0731
Ga, ppm | 6.0957 | 0.84565 | 4.4044 | 7.787 | 3.5587 | 8.6327 | 13.87% | 27.75% | 41.62% | 5.7909 | 6.4005
Gd, ppm | 1.4851 | 0.20844 | 1.0682 | 1.9019 | 0.85974 | 2.1104 | 14.04% | 28.07% | 42.11% | 1.4108 | 1.5593
Hf, ppm | 0.57778 | 0.11451 | 0.34877 | 0.8068 | 0.23426 | 0.92131 | 19.82% | 39.64% | 59.46% | 0.54889 | 0.60667
Hg, ppm | 1.0073 | 0.051829 | 0.90365 | 1.111 | 0.85183 | 1.1628 | 5.145% | 10.29% | 15.44% | 0.95695 | 1.0577
Ho, ppm | 0.1393 | 0.0063238 | 0.12666 | 0.15195 | 0.12033 | 0.15828 | 4.54% | 9.079% | 13.62% | 0.13234 | 0.14627
In, ppm | 3.6626 | 0.40487 | 2.8528 | 4.4723 | 2.448 | 4.8772 | 11.05% | 22.11% | 33.16% | 3.4794 | 3.8457
K, wt.% | 0.16692 | 0.020968 | 0.12498 | 0.20885 | 0.10401 | 0.22982 | 12.56% | 25.12% | 37.69% | 0.15857 | 0.17526
La, ppm | 10.5317 | 1.8664 | 6.799 | 14.2645 | 4.9326 | 16.1308 | 17.72% | 35.44% | 53.16% | 10.0051 | 11.0583
Li, ppm | 6.1564 | 0.84304 | 4.4703 | 7.8425 | 3.6273 | 8.6855 | 13.69% | 27.39% | 41.08% | 5.8486 | 6.4642
Lu, ppb | 27.3341 | 4.2858 | 18.7626 | 35.9057 | 14.4768 | 40.1914 | 15.68% | 31.36% | 47.04% | 25.9674 | 28.7008
Mg, wt.% | 0.10638 | 0.014388 | 0.077602 | 0.13516 | 0.063213 | 0.14954 | 13.53% | 27.05% | 40.58% | 0.10106 | 0.1117
Mn, ppm | 230.9176 | 8.4502 | 214.0173 | 247.818 | 205.5671 | 256.2682 | 3.659% | 7.319% | 10.98% | 219.3718 | 242.4635
Mo, ppm | 4.0186 | 0.36078 | 3.297 | 4.7401 | 2.9362 | 5.1009 | 8.978% | 17.96% | 26.93% | 3.8177 | 4.2195
Na, ppm | 461.4667 | 84.0248 | 293.4171 | 629.5162 | 209.3923 | 713.541 | 18.21% | 36.42% | 54.62% | 438.3933 | 484.54
Nb, ppm | 0.38918 | 0.075254 | 0.23868 | 0.53969 | 0.16342 | 0.61495 | 19.34% | 38.67% | 58.01% | 0.36972 | 0.40864
Nd, ppm | 10.5923 | 1.4128 | 7.7667 | 13.418 | 6.3539 | 14.8308 | 13.34% | 26.68% | 40.01% | 10.0627 | 11.122
Ni, ppm | 649.9844 | 27.6666 | 594.6512 | 705.3177 | 566.9846 | 732.9843 | 4.257% | 8.513% | 12.77% | 617.4852 | 682.4837
P, ppm | 247.7249 | 31.0927 | 185.5395 | 309.9104 | 154.4468 | 341.0031 | 12.55% | 25.1% | 37.65% | 235.3387 | 260.1112
Pb, ppm | 703.095 | 36.5767 | 629.9417 | 776.2483 | 593.365 | 812.8249 | 5.202% | 10.4% | 15.61% | 667.9402 | 738.2497
Pr, ppm | 2.8257 | 0.42958 | 1.9665 | 3.6848 | 1.5369 | 4.1144 | 15.2% | 30.41% | 45.61% | 2.6844 | 2.967
Rb, ppm | 8.3917 | 1.3912 | 5.6094 | 11.1741 | 4.2182 | 12.5653 | 16.58% | 33.16% | 49.73% | 7.9722 | 8.8113
S, wt.% | 4.258 | 0.21269 | 3.8326 | 4.6833 | 3.6199 | 4.896 | 4.995% | 9.99% | 14.99% | 4.0451 | 4.4709
Sb, ppm | 126.2324 | 21.9906 | 82.2512 | 170.2136 | 60.2606 | 192.2042 | 17.42% | 34.84% | 52.26% | 119.9208 | 132.544
Sc, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Se, ppm | 45.0449 | 5.0859 | 34.873 | 55.2168 | 29.7871 | 60.3027 | 11.29% | 22.58% | 33.87% | 42.7926 | 47.2971
Sm, ppm | 1.9443 | 0.2494 | 1.4455 | 2.4431 | 1.1961 | 2.6925 | 12.83% | 25.65% | 38.48% | 1.8471 | 2.0416
Sn, ppm | 2.8083 | 0.25115 | 2.306 | 3.3106 | 2.0549 | 3.5617 | 8.943% | 17.89% | 26.83% | 2.6679 | 2.9487
Sr, ppm | 37.7147 | 6.0293 | 25.6561 | 49.7734 | 19.6267 | 55.8028 | 15.99% | 31.97% | 47.96% | 35.829 | 39.6005
Tb, ppm | 0.18568 | 0.032584 | 0.12051 | 0.25085 | 0.087928 | 0.28343 | 17.55% | 35.1% | 52.65% | 0.1764 | 0.19497
Te, ppm | 25.3294 | 1.84 | 21.6493 | 29.0095 | 19.8093 | 30.8495 | 7.264% | 14.53% | 21.79% | 24.0629 | 26.5959
Th, ppm | 3.7373 | 0.70705 | 2.3232 | 5.1514 | 1.6162 | 5.8585 | 18.92% | 37.84% | 56.76% | 3.5505 | 3.9242
Ti, ppm | 97.0729 | 19.7967 | 57.4795 | 136.6662 | 37.6828 | 156.4629 | 20.39% | 40.79% | 61.18% | 92.2192 | 101.9265
Tl, ppm | 7.2658 | 1.1582 | 4.9494 | 9.5822 | 3.7912 | 10.7404 | 15.94% | 31.88% | 47.82% | 6.9025 | 7.6291
U, ppm | 1.215 | 0.16891 | 0.87721 | 1.5528 | 0.70831 | 1.7218 | 13.9% | 27.8% | 41.7% | 1.1543 | 1.2758
V, ppm | 10.1531 | 1.5518 | 7.0495 | 13.2568 | 5.4976 | 14.8086 | 15.28% | 30.57% | 45.85% | 9.6455 | 10.6608
W, ppm | 3.497 | 0.70911 | 2.0788 | 4.9152 | 1.3697 | 5.6243 | 20.28% | 40.56% | 60.83% | 3.3222 | 3.6719
Y, ppm | 3.447 | 0.50558 | 2.4359 | 4.4582 | 1.9303 | 4.9638 | 14.67% | 29.33% | 44% | 3.2747 | 3.6194
Yb, ppm | 0.19529 | 0.018443 | 0.15841 | 0.23218 | 0.13996 | 0.25062 | 9.444% | 18.89% | 28.33% | 0.18553 | 0.20506
Zn, wt.% | 0.2536 | 0.011648 | 0.23031 | 0.2769 | 0.21866 | 0.28855 | 4.593% | 9.186% | 13.78% | 0.24092 | 0.26628
Zr, ppm | 21.1335 | 2.9402 | 15.2532 | 27.0139 | 12.313 | 29.9541 | 13.91% | 27.82% | 41.74% | 20.0769 | 22.1902
Infrared Combustion
S, wt.% | 4.8539 | 0.21305 | 4.4278 | 5.28 | 4.2147 | 5.4931 | 4.389% | 8.779% | 13.17% | 4.6112 | 5.0966

## Fire Assay
| Analytical results for Ag in OREAS 604 (Certified Value 488 ppm) | [29 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 03 | 05 | 07 | 08 | 10 | 11 | 12 | 16 | 17 | 19 | 20 | 23 | [15 empty] | ppm
|  |  | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*OES | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | [15 empty] | 0
|  |  | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 40g | 30g | [15 empty] | 0
| 1 | 1 | 421.8 | 472 | 511 | 485 | 503 | 501 | 481 | 521.9 | 475 | 431 | 487 | 492 | 487 | [15 empty] | 1
| 1 | 2 | 412.3 | 475 | 499 | 480 | 520 | 491 | 480 | 476.4 | 476 | 456 | 495 | 496.5 | 486 | [15 empty] | 0
| 1 | 3 | 436.3 | 466 | 495 | 477 | 530 | 497 | 486 | 459.3 | 471 | 428 | 491 | 486.5 | 488 | [15 empty] | 16
| 1 | 4 | 431.8 | 473 | 474 | 479 | 512 | 500 | 483 | 495.7 | 474 | 442 | 491 | 495 | 486 | [15 empty] | 488.4652
| 1 | 5 | 442.8 | 472 | 475 | 486 | 527 | 496 | 485 | 511.5 | 478 | 445 | 488 | 507 | 485 | [15 empty] | 7
| 1 | 6 | 418 | 468 | 481 | 484 | 490 | 497 | 483 | 480.4 | 472 | 451 | 502 | 489.5 | 481
| Mean |  | 427.1667 | 471 | 489.1667 | 481.8333 | 513.6667 | 497 | 483 | 490.8667 | 474.3333 | 442.1667 | 492.3333 | 494.4167 | 485.5
| Median |  | 426.8 | 472 | 488 | 482 | 516 | 497 | 483 | 488.05 | 474.5 | 443.5 | 491 | 493.5 | 486
| Std Dev. |  | 11.6887 | 3.3466 | 14.8649 | 3.656 | 15.2403 | 3.5214 | 2.2804 | 23.3557 | 2.582 | 10.9803 | 5.5015 | 7.1513 | 2.429
| Rel.Std.Dev. |  | 2.736% | 0.7105% | 3.039% | 0.7588% | 2.967% | 0.7085% | 0.4721% | 4.758% | 0.5443% | 2.483% | 1.117% | 1.446% | 0.5003%
| PDM3 |  | -12.55% | -3.576% | 0.1436% | -1.358% | 5.159% | 1.747% | -1.119% | 0.4916% | -2.893% | -9.478% | 0.7919% | 1.218% | -0.607%
| Z-Score (Absolute) |  | 4.41 | 1.1 | 0.28 | 0.28 | 2.13 | 0.87 | 0.19 | 0.41 | 0.84 | 3.28 | 0.52 | 0.67 | 0
| Analytical results for Au in OREAS 604 (Certified Value 1.43 ppm) | [29 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | 2
|  |  | 1.0g | 40g | 25g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 30g | 20g | 40g | 40g | 20g | 3
| 1 | 1 | 1.41 | 1.413 | 1.43 | 1.52 | 1.4 | 1.46 | 1.44 | 1.48 | 1.28 | 1.4147 | 1.03 | 1.419 | 1.442 | 1.358 | 1.4 | 1.35 | 1.401 | 1.4 | 1.52 | 1.478 | 1.5 | 1.355 | 1.38 | 1.393 | 1.52 | 1.45 | 1.54 | 1.475 | 1
| 1 | 2 | 1.42 | 1.442 | 1.43 | 1.48 | 1.39 | 1.2 | 1.47 | 1.46 | 1.265 | 1.4127 | 1.03 | 1.392 | 1.442 | 1.357 | 1.41 | 1.5 | 1.423 | 1.49 | 1.52 | 1.516 | 1.45 | 1.343 | 1.37 | 1.431 | 1.52 | 1.44 | 1.5 | 1.44 | 0
| 1 | 3 | 1.42 | 1.428 | 1.43 | 1.47 | 1.365 | 1.445 | 1.49 | 1.475 | 1.395 | 1.3995 | 1.03 | 1.402 | 1.425 | 1.315 | 1.43 | 1.5 | 1.385 | 1.44 | 1.51 | 1.462 | 1.4 | 1.252 |  | 1.396 | 1.54 | 1.42 | 1.5 | 1.44 | 16
| 1 | 4 | 1.39 | 1.351 | 1.41 | 1.455 | 1.38 | 1.41 | 1.43 | 1.45 | 1.385 | 1.3953 | 1.03 | 1.408 | 1.47 | 1.302 | 1.4 | 1.4 | 1.411 | 1.45 | 1.5 | 1.47 | 1.3 | 1.387 |  | 1.385 | 1.48 | 1.44 | 1.5 | 1.485 | 1.427
| 1 | 5 | 1.35 | 1.367 | 1.43 | 1.475 | 1.375 | 1.42 | 1.405 | 1.555 | 1.365 | 1.4028 | 1.06 | 1.401 | 1.45 | 1.3 | 1.39 | 1.49 | 1.383 | 1.45 | 1.52 | 1.466 | 1.4 | 1.386 |  | 1.373 | 1.34 | 1.45 | 1.5 | 1.45 | 8
| 1 | 6 | 1.39 | 1.382 | 1.43 | 1.48 | 1.36 | 1.45 | 1.44 | 1.455 | 1.42 | 1.3892 | 1.03 | 1.419 | 1.488 | 1.314 | 1.41 | 1.41 | 1.426 | 1.39 | 1.5 | 1.473 | 1.45 | 1.34 |  | 1.376 | 1.54 | 1.41 | 1.51 | 1.44
|  | 7 | 00: 1.36
|  | 8 | 00: 1.38
|  | 9 | 00: 1.39
|  | 10 | 00: 1.4
|  | 11 | 00: 1.34
|  | 12 | 00: 1.34
|  | 13 | 00: 1.36
|  | 14 | 00: 1.37
|  | 15 | 00: 1.37
|  | 16 | 00: 1.41
|  | 17 | 00: 1.42
|  | 18 | 00: 1.41
|  | 19 | 00: 1.43
|  | 20 | 00: 1.39
| Mean |  | 1.3875 | 1.3972 | 1.4267 | 1.48 | 1.3783 | 1.3975 | 1.4458 | 1.4792 | 1.3517 | 1.4024 | 1.035 | 1.4068 | 1.4528 | 1.3243 | 1.4067 | 1.4417 | 1.4048 | 1.4367 | 1.5117 | 1.4775 | 1.4167 | 1.3438 | 1.375 | 1.3923 | 1.49 | 1.435 | 1.5083 | 1.455
| Median |  | 1.39 | 1.3975 | 1.43 | 1.4775 | 1.3775 | 1.4325 | 1.44 | 1.4675 | 1.375 | 1.4012 | 1.03 | 1.405 | 1.446 | 1.3145 | 1.405 | 1.45 | 1.406 | 1.445 | 1.515 | 1.4715 | 1.425 | 1.349 | 1.375 | 1.389 | 1.52 | 1.44 | 1.5 | 1.445
| Std Dev. |  | 0.027886 | 0.036008 | 0.008165 | 0.021679 | 0.015055 | 0.098577 | 0.030069 | 0.038912 | 0.064005 | 0.0098947 | 0.012247 | 0.010722 | 0.022578 | 0.026402 | 0.013663 | 0.063692 | 0.018444 | 0.036697 | 0.0098319 | 0.019655 | 0.068313 | 0.049426 | 0.0070711 | 0.020992 | 0.076681 | 0.016432 | 0.016021 | 0.02
| Rel.Std.Dev. |  | 2.01% | 2.577% | 0.5723% | 1.465% | 1.092% | 7.054% | 2.08% | 2.631% | 4.735% | 0.7056% | 1.183% | 0.7622% | 1.554% | 1.994% | 0.9713% | 4.418% | 1.313% | 2.554% | 0.6504% | 1.33% | 4.822% | 3.678% | 0.5143% | 1.508% | 5.146% | 1.145% | 1.062% | 1.375%
| PDM3 |  | -2.77% | -2.093% | -0.02568% | 3.712% | -3.413% | -2.07% | 1.317% | 3.653% | -5.281% | -1.729% | -27.47% | -1.416% | 1.808% | -7.197% | -1.427% | 1.025% | -1.556% | 0.6751% | 5.931% | 3.536% | -0.7264% | -5.83% | -3.646% | -2.432% | 4.412% | 0.5583% | 5.697% | 1.96%
| Z-Score (Absolute) |  | NA | 0.67 | 0 | 1.22 | 1.1 | 0.24 | 0.44 | 0.85 | 1.71 | 0.56 | 8.95 | 0.45 | 0.6 | 2.34 | 0.46 | 0.34 | 0.5 | 0.23 | 1.94 | 0.99 | 0.23 | 1.47 | 1.18 | 0.78 | 2.13 | 0.19 | 1.87 | 0.65
| Analytical results for Pd in OREAS 604 (Indicative Value 5.19 ppm) | [29 empty] | Indicative
Pd | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 22; 01: 25 | ppm
|  |  | 00: FA*OES; 01: FA*OES | 2
|  |  | 00: 30g; 01: 40g | 2
| 1 | 1 | 00: 5.11; 01: 5.36 | 1
| 1 | 2 | 00: 5.02; 01: 5.35 | 1
| 1 | 3 | 01: 5.23 | 16
| 1 | 4 | 01: 5.29 | 5.1917
| 1 | 5 | 01: 5.41 | 7
| 1 | 6 | 01: 5.27
| Mean |  | 00: 5.065; 01: 5.3183
| Median |  | 00: 5.065; 01: 5.32
| Std Dev. |  | 00: 0.06364; 01: 0.066458
| Rel.Std.Dev. |  | 00: 1.256%; 01: 1.25%
| PDM3 |  | 00: -2.44%; 01: 2.44%
| Z-Score (Absolute) |  | 00: 0.67; 01: 0.67
| Analytical results for Pt in OREAS 604 (Indicative Value 895 ppb) | [29 empty] | Indicative
Pt | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 22; 01: 25 | ppb
|  |  | 00: FA*OES; 01: FA*OES | 0
|  |  | 00: 30g; 01: 40g | 0
| 1 | 1 | 00: 870; 01: 940 | 1
| 1 | 2 | 00: 860; 01: 940 | 1
| 1 | 3 | 01: 915 | 16
| 1 | 4 | 01: 930 | 895.4167
| 1 | 5 | 01: 915 | 7
| 1 | 6 | 01: 915
| Mean |  | 00: 865; 01: 925.8333
| Median |  | 00: 865; 01: 922.5
| Std Dev. |  | 00: 7.0711; 01: 12.4164
| Rel.Std.Dev. |  | 00: 0.8175%; 01: 1.341%
| PDM3 |  | 00: -3.397%; 01: 3.397%
| Z-Score (Absolute) |  | 00: 0.67; 01: 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on D21: The INAA data is omitted from determination of the certified value for Au and is used solely for the calculation of Tolerance Limits and homogeneity evaluation of OREAS 600.

## Aqua Regia
| Analytical results for Ag in OREAS 604 (Certified Value 492 ppm) | [23 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | ppm
|  |  | AR*OES | AR*MS | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*OES | AR*AAS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*AAS | AR*OES | AR*MS | 0
|  |  | 0.5g | 40g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 20g | 01g | 0.5g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 01g | 40g | 40g | 0
| 1 | 1 | 504.746 | > 250 | 490 | 498 | 499 | 488 | 512 | 527.34 | 491 | 487 | > 100 | 500.2479 | > 200 | 495.6 | > 100 | - | 380 | > 100 | 452.604 | 480 | N.A. | 29.4 | 1
| 1 | 2 | 499.56 | > 250 | 482 | 506 | 493 | 491 | 521 | 494.435 | 503 | 475 | > 100 | 493.6139 | > 200 | 496.4 | > 100 | - | 372 | > 100 | 463.6786 | 479 | N.A. | 30.1 | 0
| 1 | 3 | 508.0667 | > 250 | 482 | 497 | 516 | 488 | 528 | 521.788 | 481 | 491 | > 100 | 495.7027 | > 200 | 499.5 | > 100 | - | 379 | > 100 | 465.9187 | 489 | N.A. | 29.9 | 16
| 1 | 4 | 503.484 | > 250 | 483 | 498 | 478 | 493 | 503 | 519.428 | 469 | 485 | > 100 | 491.9612 | > 200 | 503.8 | > 100 | - | 383 | > 100 | 463.7054 | 480 | N.A. | 28.7 | 492.363
| 1 | 5 | 506.8433 | > 250 | 478 | 499 | 474 | 496 | 509 | 499.278 | 481 | 479 | > 100 | 500.0992 | > 200 | 484.9 | > 100 | - | 342 | > 100 | 460.1995 | 478 | N.A. | 28.5 | 69
| 1 | 6 | 498.7533 | > 250 | 484 | 499 | 507 | 492 | 512 | 502.69 | 484 | 493 | > 100 | 495.2173 | > 200 | 488.9 | > 100 | - | 351 | > 100 | 469.8566 | 475 | N.A. | 24.8
| Mean |  | 503.5756 |  | 483.1667 | 499.5 | 494.5 | 491.3333 | 514.1667 | 510.8265 | 484.8333 | 485 |  | 496.1404 |  | 494.85 |  |  | 367.8333 |  | 462.6605 | 480.1667 |  | 28.5667
| Median |  | 504.115 |  | 482.5 | 498.5 | 496 | 491.5 | 512 | 511.059 | 482.5 | 486 |  | 495.46 |  | 496 |  |  | 375.5 |  | 463.692 | 479.5 |  | 29.05
| Std Dev. |  | 3.7843 |  | 3.92 | 3.2711 | 16.331 | 3.0768 | 8.9312 | 13.6754 | 11.3915 | 6.9282 |  | 3.3893 |  | 6.9151 |  |  | 17.1513 |  | 5.8603 | 4.7081 |  | 1.9511
| Rel.Std.Dev. |  | 0.7515% |  | 0.8113% | 0.6549% | 3.303% | 0.6262% | 1.737% | 2.677% | 2.35% | 1.428% |  | 0.6831% |  | 1.397% |  |  | 4.663% |  | 1.267% | 0.9805% |  | 6.83%
| PDM3 |  | 2.277% |  | -1.868% | 1.45% | 0.434% | -0.2091% | 4.428% | 3.75% | -1.529% | -1.495% |  | 0.7672% |  | 0.5051% |  |  | -25.29% |  | -6.033% | -2.477% |  | -94.2%
| Z-Score (Absolute) |  | 1.01 | NA | 0.67 | 0.67 | 0.26 | 0 | 1.89 | 1.61 | 0.54 | 0.52 | NA | 0.4 | NA | 0.29 | NA | NA | 10.2 | NA | 2.37 | 0.92 | NA | 38.15
| Analytical results for Al in OREAS 604 (Certified Value 0.79 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 1.085 | 1.0273 | 0.84 | 0.71 | 0.8 | 0.72 | 1.0106 | 0.56 | 0.78 | 0.84 | 0.90802 | 0.78 | 1 | 0.68 | 0.87 | 0.6 | 0.742 | 0.56 | 0.76 | 0.78 | 0.62 |  | 1
| 1 | 2 | 1.085 | 1.0091 | 0.83 | 0.69 | 0.77 | 0.75 | 0.9712 | 0.62 | 0.74 | 0.86 | 0.89438 | 0.8 | 1.07 | 0.71 | 0.83 | 0.63 | 0.785 | 0.57 | 0.76 | 0.76 | 0.61 |  | 0
| 1 | 3 | 1.085 | 1.0382 | 0.82 | 0.71 | 0.79 | 0.76 | 1.0188 | 0.59 | 0.75 | 0.85 | 0.89545 | 0.79 | 1.06 | 0.72 | 0.86 | 0.62 | 0.816 | 0.59 | 0.76 | 0.78 | 0.63 |  | 16
| 1 | 4 | 1.1 | 1.0545 | 0.82 | 0.71 | 0.8 | 0.77 | 1.0188 | 0.58 | 0.75 | 0.83 | 0.8926 | 0.77 | 1.06 | 0.7 | 0.86 | 0.64 | 0.791 | 0.59 | 0.72 | 0.85 | 0.64 |  | 0.78979
| 1 | 5 | 1.1133 | 1.0587 | 0.81 | 0.71 | 0.79 | 0.79 | 0.9636 | 0.56 | 0.75 | 0.82 | 0.9185 | 0.78 | 0.97 | 0.7 | 0.89 | 0.64 | 0.803 | 0.59 | 0.75 | 0.93 | 0.65 |  | 70
| 1 | 6 | 1.085 | 1.0464 | 0.81 | 0.76 | 0.79 | 0.77 | 1.0073 | 0.63 | 0.75 | 0.81 | 0.90918 | 0.77 | 1.03 | 0.7 | 0.88 | 0.61 | 0.795 | 0.6 | 0.77 | 0.85 | 0.66
| Mean |  | 1.0922 | 1.039 | 0.82167 | 0.715 | 0.79 | 0.76 | 0.99838 | 0.59 | 0.75333 | 0.835 | 0.90302 | 0.78167 | 1.0317 | 0.70167 | 0.865 | 0.62333 | 0.78867 | 0.58333 | 0.75333 | 0.825 | 0.635
| Median |  | 1.085 | 1.0423 | 0.82 | 0.71 | 0.79 | 0.765 | 1.009 | 0.585 | 0.75 | 0.835 | 0.90174 | 0.78 | 1.045 | 0.7 | 0.865 | 0.625 | 0.793 | 0.59 | 0.76 | 0.815 | 0.635
| Std Dev. |  | 0.011957 | 0.018518 | 0.01169 | 0.023452 | 0.010954 | 0.023664 | 0.024541 | 0.029665 | 0.013663 | 0.018708 | 0.010422 | 0.01169 | 0.039707 | 0.013292 | 0.020736 | 0.01633 | 0.025256 | 0.015055 | 0.017512 | 0.064109 | 0.018708
| Rel.Std.Dev. |  | 1.095% | 1.782% | 1.423% | 3.28% | 1.387% | 3.114% | 2.458% | 5.028% | 1.814% | 2.241% | 1.154% | 1.496% | 3.849% | 1.894% | 2.397% | 2.62% | 3.202% | 2.581% | 2.325% | 7.771% | 2.946%
| PDM3 |  | 38.29% | 31.56% | 4.036% | -9.469% | 0.02677% | -3.772% | 26.41% | -25.3% | -4.616% | 5.725% | 14.34% | -1.028% | 30.63% | -11.16% | 9.523% | -21.08% | -0.142% | -26.14% | -4.616% | 4.458% | -19.6%
| Z-Score (Absolute) |  | 2.43 | 2 | 0.25 | 0.67 | 0 | 0.24 | 1.67 | 1.61 | 0.34 | 0.36 | 0.91 | 0.07 | 1.94 | 0.71 | 0.6 | 1.34 | 0.06 | 1.66 | 0.24 | 0.28 | 1.24
| Analytical results for As in OREAS 604 (Certified Value 966 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 0
| 1 | 1 | 1060.545 | 942 | 1010 | 994 | 1020 | 967 | 1004.018 | 990 | 964 | 886 | 1163.6968 | 980 | 1054 | 899 | 785 | 911 | 910.849 | 918 | 976 | 912 | 950 |  | 1
| 1 | 2 | 1060.765 | 936 | 988 | 965 | 1025 | 989 | 1009.7037 | 990 | 953 | 921 | 1126.5939 | 990 | 1043 | 904 | 650 | 896 | 919.113 | 950 | 973 | 920 | 998 |  | 0
| 1 | 3 | 1073.44 | 931 | 1010 | 954 | 1035 | 991 | 1006.777 | 1010 | 952 | 902 | 1109.8811 | 980 | 1053 | 894 | 770 | 874 | 918.66 | 949 | 980 | 896 | 979 |  | 16
| 1 | 4 | 1091.476 | 931 | 1000 | 970 | 1025 | 982 | 983.5898 | 960 | 959 | 887 | 1128.7346 | 961 | 1066 | 912 | 775 | 904 | 897.041 | 949 | 943 | 880 | 960 |  | 966.0952
| 1 | 5 | 1077.412 | 942 | 1010 | 954 | 1035 | 1005 | 995.968 | 980 | 946 | 889 | 1134.0605 | 992 | 1021 | 920 | 745 | 873 | 914.551 | 939 | 970 | 890 | 950 |  | 71
| 1 | 6 | 1085.044 | 938 | 1010 | 1030 | 1030 | 1000 | 991.4929 | 980 | 966 | 896 | 1134.5833 | 956 | 1017 | 886 | 730 | 886 | 907.006 | 969 | 991 | 912 | 979
| Mean |  | 1074.7803 | 936.6667 | 1004.6667 | 977.8333 | 1028.3333 | 989 | 998.5916 | 985 | 956.6667 | 896.8333 | 1132.925 | 976.5 | 1042.3333 | 902.5 | 742.5 | 890.6667 | 911.2033 | 945.6667 | 972.1667 | 901.6667 | 969.3333
| Median |  | 1075.426 | 937 | 1010 | 967.5 | 1027.5 | 990 | 999.993 | 985 | 956 | 892.5 | 1131.3976 | 980 | 1048 | 901.5 | 757.5 | 891 | 912.7 | 949 | 974.5 | 904 | 969.5
| Std Dev. |  | 12.5848 | 4.9666 | 9.0921 | 29.4782 | 6.0553 | 13.5204 | 10.0225 | 16.4317 | 7.6855 | 13.3179 | 17.5477 | 14.8829 | 19.5312 | 12.2923 | 49.6739 | 15.6929 | 8.3389 | 16.7053 | 16.0427 | 15.41 | 19.2007
| Rel.Std.Dev. |  | 1.171% | 0.5302% | 0.905% | 3.015% | 0.5888% | 1.367% | 1.004% | 1.668% | 0.8034% | 1.485% | 1.549% | 1.524% | 1.874% | 1.362% | 6.69% | 1.762% | 0.9152% | 1.767% | 1.65% | 1.709% | 1.981%
| PDM3 |  | 11.25% | -3.046% | 3.993% | 1.215% | 6.442% | 2.371% | 3.364% | 1.957% | -0.9759% | -7.169% | 17.27% | 1.077% | 7.891% | -6.583% | -23.14% | -7.808% | -5.682% | -2.115% | 0.6285% | -6.669% | 0.3352%
| Z-Score (Absolute) |  | 2.01 | 0.62 | 0.67 | 0.04 | 1.13 | 0.38 | 0.56 | 0.3 | 0.24 | 1.38 | 3.12 | 0.14 | 1.39 | 1.28 | 3.98 | 1.5 | 1.11 | 0.45 | 0.17 | 1.29 | 0
| Analytical results for Au in OREAS 604 (Certified Value 1.43 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 00 | 01 | 02 | 05 | 06 | 07 | 09 | 14 | 15 | 18 | 24 | 25 | 26 | 27 | [8 empty] | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | [8 empty] | 2
|  |  | 1.0g | 30g | 25g | 25g | 25g | 25g | 25g | 25g | 30g | 50g | 40g | 40g | 40g | 40g | [8 empty] | 3
| 1 | 1 | 1.41 | 1.411 | 1.382 | 1.45 | 1.27 | 1.5 | 1.4337 | 1.317 | 1.52 | 1.429 | 1.49 | 1.49 | 1.47 | 1.4785 | [8 empty] | 1
| 1 | 2 | 1.42 | 1.423 | 1.428 | 1.43 | 1.25 | 1.45 | 1.4365 | 1.316 | 1.46 | 1.428 | 1.42 | 1.48 | 1.48 | 1.314 | [8 empty] | 0
| 1 | 3 | 1.42 | 1.343 | 1.41 | 1.42 | 1.33 | 1.48 | 1.4053 | 1.329 | 1.44 | 1.436 | 1.49 | 1.48 | 1.47 | 1.453 | [8 empty] | 16
| 1 | 4 | 1.39 | 1.432 | 1.371 | 1.42 | 1.1 | 1.44 | 1.3799 | 1.289 | 1.51 | 1.466 | 1.44 | 1.47 | 1.46 | 1.383 | [8 empty] | 1.428
| 1 | 5 | 1.35 | 1.379 | 1.353 | 1.41 | 1.37 | 1.5 | 1.3484 | 1.249 | 1.52 | 1.424 | 1.44 | 1.45 | 1.49 | 1.4695 | [8 empty] | 72
| 1 | 6 | 1.39 | 1.37 | 1.368 | 1.4 | 1.23 | 1.5 | 1.3705 | 1.304 | 1.53 | 1.452 | 1.45 | 1.45 | 1.45 | 1.423
|  | 7 | 01: 1.36
|  | 8 | 01: 1.38
|  | 9 | 01: 1.39
|  | 10 | 01: 1.4
|  | 11 | 01: 1.34
|  | 12 | 01: 1.34
|  | 13 | 01: 1.36
|  | 14 | 01: 1.37
|  | 15 | 01: 1.37
|  | 16 | 01: 1.41
|  | 17 | 01: 1.42
|  | 18 | 01: 1.41
|  | 19 | 01: 1.43
|  | 20 | 01: 1.39
| Mean |  | 1.3875 | 1.393 | 1.3853 | 1.4217 | 1.2583 | 1.4783 | 1.3957 | 1.3007 | 1.4967 | 1.4392 | 1.455 | 1.47 | 1.47 | 1.4202
| Median |  | 1.39 | 1.395 | 1.3765 | 1.42 | 1.26 | 1.49 | 1.3926 | 1.31 | 1.515 | 1.4325 | 1.445 | 1.475 | 1.47 | 1.438
| Std Dev. |  | 0.027886 | 0.034554 | 0.028254 | 0.017224 | 0.09347 | 0.027142 | 0.035544 | 0.02871 | 0.037238 | 0.016449 | 0.02881 | 0.016733 | 0.014142 | 0.06257
| Rel.Std.Dev. |  | 2.01% | 2.481% | 2.039% | 1.212% | 7.428% | 1.836% | 2.547% | 2.207% | 2.488% | 1.143% | 1.98% | 1.138% | 0.9621% | 4.406%
| PDM3 |  | -2.836% | -2.451% | -2.988% | -0.4439% | -11.88% | 3.524% | -2.26% | -8.917% | 4.808% | 0.7816% | 1.89% | 2.941% | 2.941% | -0.5489%
| Z-Score (Absolute) |  | NA | 0.53 | 0.67 | 0 | 3.03 | 1.05 | 0.48 | 2.05 | 1.39 | 0.32 | 0.62 | 0.9 | 0.9 | 0.03
| Analytical results for B in OREAS 604 (Certified Value < 10 ppm) | [23 empty] | Certified
B | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 05; 02: 06; 05: 07; 06: 08; 07: 09; 08: 14; 09: 18; 10: 19; 11: 21; 12: 23 |  | ppm
|  |  | 01: AR*OES; 02: AR*OES; 05: AR*OES; 06: AR*OES; 07: AR*OES/MS; 08: AR*OES; 09: AR*OES; 10: AR*OES/MS; 11: AR*OES; 12: AR*OES |  | 1
|  |  | 01: 01g; 02: 01g; 05: 01g; 06: 01g; 07: 20g; 08: 0.5g; 09: 10g; 10: 0.25g; 11: 0.25g; 12: 0.5g |  | 1
| 1 | 1 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 19.802; 11: < 10; 12: < 20 |  | 1
| 1 | 2 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 19.668; 11: < 10; 12: < 20 |  | 3
| 1 | 3 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 20.75; 11: < 10; 12: < 20 |  | 16
| 1 | 4 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 20.026; 11: < 10; 12: < 20 |  | < 10
| 1 | 5 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 20.315; 11: < 10; 12: < 20 |  | 73
| 1 | 6 | 01: < 10; 02: < 10; 05: < 10; 06: < 10; 07: < 10; 08: < 5; 09: < 10; 10: 19.946; 11: < 10; 12: < 20
| Mean |  | 10: 20.0845
| Median |  | 10: 19.986
| Std Dev. |  | 10: 0.39291
| Rel.Std.Dev. |  | 10: 1.956%
PDM3
Z-Score (Absolute) | bar: 01 to 12
| Analytical results for Ba in OREAS 604 (Indicative Value 410 ppm) | [23 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 23 | 25 | [5 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | [5 empty] | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g | [5 empty] | 0
| 1 | 1 | 10.6533 | 124 | 40 | 20 | 10 | 40 | 35 | 997 | 46 | 101.097 | 18 | 78 | 51 | 93 | 344 | 13 | 32 | [5 empty] | 1
| 1 | 2 | 10.5387 | 124 | 40 | 20 | 10 | 40 | 32 | 501 | 49 | 102.306 | 18 | 73 | 51 | 84 | 266 | 13 | 36 | [5 empty] | 4
| 1 | 3 | 10.5653 | 125 | 40 | 20 | 10 | 30 | 33.2772 | 517 | 52 | 107.188 | 18 | 102 | 49 | 87 | 338 | 11 | 32 | [5 empty] | 16
| 1 | 4 | 10.5307 | 120 | 30 | 20 | 10 | 30 | 33.2793 | 455 | 53 | 102.974 | 20 | 97 | 51 | 88 | 326 | 12 | 31 | [5 empty] | 409.5
| 1 | 5 | 10.7227 | 111 | 30 | 20 | 10 | 30 | 32.4277 | 472 | 58 | 104.7671 | 21 | 77 | 52 | 88 | 321 | 14 | 31 | [5 empty] | 12
| 1 | 6 | 10.4213 | 113 | 30 | 20 | 10 | 30 | 35.877 | 512 | 49 | 106.5896 | 18 | 97 | 50 | 91 | 309 | 13 | 32
| Mean |  | 10.572 | 119.5 | 35 | 20 | 10 | 33.3333 | 33.6435 | 575.6667 | 51.1667 | 104.1536 | 18.8333 | 87.3333 | 50.6667 | 88.5 | 317.3333 | 12.6667 | 32.3333
| Median |  | 10.552 | 122 | 35 | 20 | 10 | 30 | 33.2783 | 506.5 | 50.5 | 103.8705 | 18 | 87.5 | 51 | 88 | 323.5 | 13 | 32
| Std Dev. |  | 0.10477 | 6.091 | 5.4772 | 0 | 0 | 5.164 | 1.5017 | 207.7977 | 4.1673 | 2.4358 | 1.3292 | 12.6596 | 1.0328 | 3.1464 | 28.0404 | 1.0328 | 1.8619
| Rel.Std.Dev. |  | 0.991% | 5.097% | 15.65% | 0% | 0% | 15.49% | 4.463% | 36.1% | 8.145% | 2.339% | 7.057% | 14.5% | 2.038% | 3.555% | 8.836% | 8.154% | 5.758%
| PDM3 |  | -97.42% | -70.82% | -91.45% | -95.12% | -97.56% | -91.86% | -91.78% | 40.58% | -87.51% | -74.57% | -95.4% | -78.67% | -87.63% | -78.39% | -22.51% | -96.91% | -92.1%
| Z-Score (Absolute) |  | 0.74 | 2.55 | 0 | 0.45 | 0.75 | 0.05 | 0.04 | 13.78 | 0.49 | 2.09 | 0.5 | 1.58 | 0.47 | 1.62 | 8.83 | 0.67 | 0.1
| Analytical results for Be in OREAS 604 (Certified Value 0.33 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | [6 empty] | 2
| 1 | 1 | < 1 | 0.31 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.37842 | < 0.5 | 0.35485 | 0.36 | < 0.5 | < 0.5 | 0.3 | < 1 | 0.3 | < 2 | [6 empty] | 1
| 1 | 2 | < 1 | 0.32 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.41386 | < 0.5 | 0.35421 | 0.43 | < 0.5 | < 0.5 | 0.2 | < 1 | 0.35 | < 2 | [6 empty] | 0
| 1 | 3 | < 1 | 0.3 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.41786 | < 0.5 | 0.36609 | 0.35 | < 0.5 | < 0.5 | 0.3 | < 1 | 0.32 | < 2 | [6 empty] | 16
| 1 | 4 | < 1 | 0.31 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.42677 | < 0.5 | 0.35079 | 0.37 | < 0.5 | < 0.5 | 0.3 | < 1 | 0.26 | < 2 | [6 empty] | 0.33436
| 1 | 5 | < 1 | 0.3 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.39975 | < 0.5 | 0.38079 | 0.38 | < 0.5 | < 0.5 | 0.3 | < 1 | 0.32 | < 2 | [6 empty] | 74
| 1 | 6 | < 1 | 0.3 | < 0.5 | < 0.5 | < 0.5 | < 0.5 | 0.4299 | < 0.5 | 0.35762 | 0.37 | < 0.5 | < 0.5 | 0.2 | < 1 | 0.22 | < 2
| Mean |  | 02: 0.30667; 09: 0.41109; 13: 0.36073; 14: 0.37667; 17: 0.26667; 19: 0.295
| Median |  | 02: 0.305; 09: 0.41586; 13: 0.35623; 14: 0.37; 17: 0.3; 19: 0.31
| Std Dev. |  | 02: 0.008165; 09: 0.019232; 13: 0.011106; 14: 0.028048; 17: 0.05164; 19: 0.047223
| Rel.Std.Dev. |  | 02: 2.662%; 09: 4.678%; 13: 3.079%; 14: 7.446%; 17: 19.36%; 19: 16.01%
| PDM3 |  | 02: -8.282%; 09: 22.95%; 13: 7.886%; 14: 12.65%; 17: -20.25%; 19: -11.77%
| Z-Score (Absolute) |  | NA | 0.51 | NA | NA | NA | NA | 1.47 | NA | 0.51 | 0.61 | NA | NA | 1.27 | NA | 0.74 | NA
| Analytical results for Bi in OREAS 604 (Certified Value 26.9 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  | 1
| 1 | 1 | 39.7433 | 30.15 | 38 | 27 | 29 | 26 | 22.861 | 30 | 25 | 35 | 23.6482 | 29 | 32 | < 10 | 19.9 | 26.836 | 17 | 23 | 28.6 | 29.2 |  |  | 1
| 1 | 2 | 39.204 | 27.62 | 39 | 27 | 31 | 28 | 22.7635 | 40 | 27 | 40 | 24.446 | 28 | 29 | < 10 | 15.8 | 27.009 | 20 | 24 | 29.6 | 29 |  |  | 0
| 1 | 3 | 39.01 | 27.84 | 42 | 26 | 32 | 26 | 22.5244 | 40 | 25 | 37 | 23.0669 | 28 | 29 | < 10 | 17.9 | 27.332 | 21 | 21 | 25.6 | 28.5 |  |  | 16
| 1 | 4 | 39.545 | 27.32 | 41 | 26 | 34 | 28 | 22.8605 | 30 | 25 | 32 | 24.9475 | 28 | 27 | < 10 | 20.7 | 26.31 | 22 | 21 | 25.5 | 28.5 |  |  | 26.9191
| 1 | 5 | 38.46 | 28.2 | 39 | 28 | 34 | 28 | 23.23 | 30 | 26 | 34 | 24.582 | 29 | 31 | < 10 | 17.8 | 26.721 | 20 | 21 | 26.8 | 28.9 |  |  | 75
| 1 | 6 | 39.445 | 28.84 | 39 | 29 | 31 | 25 | 23.3945 | 30 | 28 | 36 | 24.2438 | 29 | 29 | < 10 | 16.9 | 26.7 | 21 | 24 | 26.9 | 28.2
| Mean |  | 39.2346 | 28.3283 | 39.6667 | 27.1667 | 31.8333 | 26.8333 | 22.939 | 33.3333 | 26 | 35.6667 | 24.1557 | 28.5 | 29.5 |  | 18.1667 | 26.818 | 20.1667 | 22.3333 | 27.1667 | 28.7167
| Median |  | 39.3245 | 28.02 | 39 | 27 | 31.5 | 27 | 22.8608 | 30 | 25.5 | 35.5 | 24.3449 | 28.5 | 29 |  | 17.85 | 26.7785 | 20.5 | 22 | 26.85 | 28.7
| Std Dev. |  | 0.45843 | 1.0348 | 1.5055 | 1.169 | 1.9408 | 1.3292 | 0.31846 | 5.164 | 1.2649 | 2.7325 | 0.68465 | 0.54772 | 1.7607 |  | 1.8348 | 0.34129 | 1.7224 | 1.5055 | 1.6379 | 0.37639
| Rel.Std.Dev. |  | 1.168% | 3.653% | 3.795% | 4.303% | 6.097% | 4.953% | 1.388% | 15.49% | 4.865% | 7.661% | 2.834% | 1.922% | 5.968% |  | 10.1% | 1.273% | 8.541% | 6.741% | 6.029% | 1.311%
| PDM3 |  | 45.75% | 5.235% | 47.36% | 0.9197% | 18.26% | -0.3185% | -14.79% | 23.83% | -3.414% | 32.5% | -10.27% | 5.873% | 9.588% |  | -32.51% | -0.3755% | -25.08% | -17.04% | 0.9197% | 6.678%
| Z-Score (Absolute) |  | 2.7 | 0.26 | 2.8 | 0 | 1.05 | 0.07 | 0.95 | 1.38 | 0.26 | 1.9 | 0.67 | 0.3 | 0.52 | NA | 2.02 | 0.08 | 1.57 | 1.08 | 0 | 0.35
| Analytical results for Ca in OREAS 604 (Certified Value 0.591 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 0.53 | 0.6 | 0.6 | 0.6 | 0.62 | 0.6 | 0.64491 | 0.58 | 0.65 | 0.63 | 0.51843 | 0.59 | 0.66 | 0.55 | 0.61 | 0.55 | 0.568 | 0.56 | 0.58 | 0.58 | 0.61 |  | 1
| 1 | 2 | 0.53 | 0.59 | 0.6 | 0.59 | 0.63 | 0.62 | 0.65408 | 0.59 | 0.63 | 0.63 | 0.51345 | 0.6 | 0.66 | 0.56 | 0.56 | 0.55 | 0.581 | 0.57 | 0.58 | 0.56 | 0.61 |  | 0
| 1 | 3 | 0.52667 | 0.59 | 0.59 | 0.59 | 0.61 | 0.6 | 0.62961 | 0.6 | 0.64 | 0.63 | 0.51899 | 0.59 | 0.67 | 0.56 | 0.59 | 0.54 | 0.608 | 0.57 | 0.58 | 0.58 | 0.62 |  | 16
| 1 | 4 | 0.52333 | 0.59 | 0.6 | 0.59 | 0.62 | 0.61 | 0.64255 | 0.57 | 0.65 | 0.63 | 0.51173 | 0.6 | 0.67 | 0.55 | 0.59 | 0.55 | 0.588 | 0.58 | 0.57 | 0.59 | 0.62 |  | 0.59112
| 1 | 5 | 0.51 | 0.6 | 0.61 | 0.59 | 0.61 | 0.63 | 0.6114 | 0.58 | 0.64 | 0.61 | 0.52008 | 0.6 | 0.64 | 0.56 | 0.59 | 0.55 | 0.598 | 0.57 | 0.57 | 0.63 | 0.62 |  | 76
| 1 | 6 | 0.50667 | 0.59 | 0.6 | 0.63 | 0.62 | 0.61 | 0.63394 | 0.58 | 0.63 | 0.64 | 0.51472 | 0.6 | 0.65 | 0.56 | 0.58 | 0.53 | 0.59 | 0.58 | 0.59 | 0.57 | 0.61
| Mean |  | 0.52111 | 0.59333 | 0.6 | 0.59833 | 0.61833 | 0.61167 | 0.63608 | 0.58333 | 0.64 | 0.62833 | 0.51624 | 0.59667 | 0.65833 | 0.55667 | 0.58667 | 0.545 | 0.58883 | 0.57167 | 0.57833 | 0.585 | 0.615
| Median |  | 0.525 | 0.59 | 0.6 | 0.59 | 0.62 | 0.61 | 0.63825 | 0.58 | 0.64 | 0.63 | 0.51658 | 0.6 | 0.66 | 0.56 | 0.59 | 0.55 | 0.589 | 0.57 | 0.58 | 0.58 | 0.615
| Std Dev. |  | 0.010256 | 0.005164 | 0.0063246 | 0.016021 | 0.0075277 | 0.01169 | 0.014828 | 0.010328 | 0.0089443 | 0.0098319 | 0.0033939 | 0.005164 | 0.01169 | 0.005164 | 0.01633 | 0.0083666 | 0.013776 | 0.0075277 | 0.0075277 | 0.02429 | 0.0054772
| Rel.Std.Dev. |  | 1.968% | 0.8703% | 1.054% | 2.678% | 1.217% | 1.911% | 2.331% | 1.771% | 1.398% | 1.565% | 0.6574% | 0.8655% | 1.776% | 0.9277% | 2.784% | 1.535% | 2.339% | 1.317% | 1.302% | 4.152% | 0.8906%
| PDM3 |  | -11.84% | 0.3738% | 1.502% | 1.22% | 4.603% | 3.475% | 7.606% | -1.318% | 8.268% | 6.295% | -12.67% | 0.9377% | 11.37% | -5.829% | -0.754% | -7.803% | -0.3875% | -3.292% | -2.164% | -1.036% | 4.039%
| Z-Score (Absolute) |  | 2.35 | 0.04 | 0.27 | 0 | 0.87 | 0.65 | 1.46 | 0.29 | 1.59 | 1.2 | 2.51 | 0.15 | 2.2 | 1.17 | 0.18 | 1.56 | 0.11 | 0.67 | 0.45 | 0.53 | 0.76
| Analytical results for Cd in OREAS 604 (Certified Value 15.5 ppm) | [23 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 2
| 1 | 1 | 23.8105 | 13.97 | 15.5 | 15.6 | 15.6 | 15.2 | 17.0586 | 17 | 16 | 16.3 | 14.4529 | 16 | 15.2 | 14 | 14.1 | 10 | 13.623 | 16 | 14.9 | 18.8 | 15 |  | 1
| 1 | 2 | 23.7389 | 14.13 | 15.5 | 15 | 15.8 | 15.4 | 17.893 | 18 | 16 | 16.8 | 13.7825 | 16.1 | 15.4 | 14 | 11.6 | 10 | 13.87 | 17 | 15 | 19.1 | 15.3 |  | 0
| 1 | 3 | 23.8908 | 14.19 | 15.4 | 14.9 | 16.1 | 15.4 | 17.8904 | 18 | 16 | 16.6 | 14.758 | 15.8 | 15.4 | 14 | 14.5 | 10 | 13.502 | 17 | 15.3 | 17.2 | 15.3 |  | 16
| 1 | 4 | 23.5497 | 14.21 | 15.6 | 15.3 | 16.1 | 15.4 | 17.7213 | 17 | 16 | 16.2 | 13.7865 | 15.6 | 15.3 | 14 | 15.2 | 9 | 13.363 | 16 | 14.6 | 16.9 | 14.9 |  | 15.5259
| 1 | 5 | 23.5382 | 14.17 | 15.6 | 15.2 | 15.9 | 15.7 | 17.2353 | 18 | 16 | 16.4 | 14.0213 | 16.2 | 14.7 | 14 | 13 | 9 | 13.437 | 16 | 14.9 | 18 | 15.2 |  | 77
| 1 | 6 | 22.8731 | 14.54 | 15.4 | 16 | 16.1 | 15.5 | 17.877 | 18 | 16 | 16.5 | 14.246 | 15.6 | 14.9 | 14 | 13.3 | 9 | 13.222 | 17 | 15.2 | 17.5 | 15
| Mean |  | 23.5669 | 14.2017 | 15.5 | 15.3333 | 15.9333 | 15.4333 | 17.6126 | 17.6667 | 16 | 16.4667 | 14.1745 | 15.8833 | 15.15 | 14 | 13.6167 | 9.5 | 13.5028 | 16.5 | 14.9833 | 17.9167 | 15.1167
| Median |  | 23.6443 | 14.18 | 15.5 | 15.25 | 16 | 15.4 | 17.7991 | 18 | 16 | 16.45 | 14.1336 | 15.9 | 15.25 | 14 | 13.7 | 9.5 | 13.4695 | 16.5 | 14.95 | 17.75 | 15.1
| Std Dev. |  | 0.36777 | 0.18681 | 0.089443 | 0.40825 | 0.20656 | 0.1633 | 0.37061 | 0.5164 | 0 | 0.21602 | 0.38748 | 0.25626 | 0.2881 | 0 | 1.2703 | 0.54772 | 0.2245 | 0.54772 | 0.24833 | 0.88412 | 0.17224
| Rel.Std.Dev. |  | 1.561% | 1.315% | 0.577% | 2.662% | 1.296% | 1.058% | 2.104% | 2.923% | 0% | 1.312% | 2.734% | 1.613% | 1.902% | 0% | 9.329% | 5.766% | 1.663% | 3.32% | 1.657% | 4.935% | 1.139%
| PDM3 |  | 51.79% | -8.529% | -0.1667% | -1.24% | 2.624% | -0.596% | 13.44% | 13.79% | 3.054% | 6.06% | -8.704% | 2.302% | -2.421% | -9.828% | -12.3% | -38.81% | -13.03% | 6.274% | -3.494% | 15.4% | -2.636%
| Z-Score (Absolute) |  | 5.23 | 0.78 | 0.04 | 0.06 | 0.32 | 0 | 1.38 | 1.41 | 0.36 | 0.65 | 0.8 | 0.28 | 0.18 | 0.91 | 1.15 | 3.75 | 1.22 | 0.67 | 0.28 | 1.57 | 0.2
| Analytical results for Ce in OREAS 604 (Certified Value 23.5 ppm) | [23 empty] | Certified
Ce | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*OES; 06: AR*MS; 07: AR*MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*OES/MS; 11: AR*MS |  | 1
|  |  | 01: 40g; 02: 20g; 05: 0.15g; 06: 01g; 07: 0.5g; 08: 01g; 09: 0.25g; 10: 0.25g; 11: 40g |  | 1
| 1 | 1 | 01: 25.16; 02: 28.1165; 05: 23; 06: 21.9534; 07: 23.19; 08: 28.6; 09: 23.1; 10: 21.416; 11: 16.9 |  | 1
| 1 | 2 | 01: 23.79; 02: 28.7273; 05: 24; 06: 21.9291; 07: 23.21; 08: 28.32; 09: 18.9; 10: 21.88; 11: 17.7 |  | 0
| 1 | 3 | 01: 23.98; 02: 27.8896; 05: 24; 06: 21.9593; 07: 23.24; 08: 29.41; 09: 23.8; 10: 22.201; 11: 17.5 |  | 16
| 1 | 4 | 01: 24.11; 02: 28.2973; 05: 23; 06: 21.2396; 07: 23.16; 08: 29.84; 09: 25.3; 10: 21.123; 11: 17.4 |  | 23.5183
| 1 | 5 | 01: 24.1; 02: 27.1432; 05: 25; 06: 21.7974; 07: 24.08; 08: 28.29; 09: 21.4; 10: 21.486; 11: 17.4 |  | 78
| 1 | 6 | 01: 24.56; 02: 28.1999; 05: 25; 06: 21.0026; 07: 23.53; 08: 28.34; 09: 22.2; 10: 21.016; 11: 17.5
| Mean |  | 01: 24.2833; 02: 28.0623; 05: 24; 06: 21.6469; 07: 23.4017; 08: 28.8; 09: 22.45; 10: 21.5203; 11: 17.4
| Median |  | 01: 24.105; 02: 28.1582; 05: 24; 06: 21.8632; 07: 23.225; 08: 28.47; 09: 22.65; 10: 21.451; 11: 17.45
| Std Dev. |  | 01: 0.49883; 02: 0.52799; 05: 0.89443; 06: 0.41828; 07: 0.35852; 08: 0.66269; 09: 2.197; 10: 0.45099; 11: 0.26833
| Rel.Std.Dev. |  | 01: 2.054%; 02: 1.881%; 05: 3.727%; 06: 1.932%; 07: 1.532%; 08: 2.301%; 09: 9.786%; 10: 2.096%; 11: 1.542%
| PDM3 |  | 01: 3.253%; 02: 19.32%; 05: 2.048%; 06: -7.957%; 07: -0.4958%; 08: 22.46%; 09: -4.542%; 10: -8.495%; 11: -26.02%
| Z-Score (Absolute) |  | 01: 0.34; 02: 1.79; 05: 0.23; 06: 0.67; 07: 0; 08: 2.07; 09: 0.37; 10: 0.72; 11: 2.27
| Analytical results for Co in OREAS 604 (Certified Value 42.6 ppm) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 2
| 1 | 1 | 38.547 | 40.6 | 43 | 42 | 44 | 42 | 41.8345 | 49 | 46 | 46 | 41.6213 | 42 | 45.76 | 40 | 40.8 | 41 | 41.03 | 44 | 45 | 41.6 | 44 |  | 1
| 1 | 2 | 38.259 | 39.8 | 44 | 41 | 45 | 43 | 42.7553 | 51 | 43 | 47 | 41.5914 | 44 | 46.2 | 40 | 33.1 | 41 | 40.98 | 44 | 44 | 42.4 | 46 |  | 0
| 1 | 3 | 38.6415 | 38.2 | 43 | 42 | 46 | 43 | 41.7178 | 52 | 42 | 46 | 41.0424 | 42 | 46.82 | 41 | 41.1 | 39 | 40.79 | 45 | 45 | 42 | 46 |  | 16
| 1 | 4 | 37.674 | 39 | 44 | 41 | 45 | 44 | 43.0823 | 49 | 43 | 45 | 41.3937 | 43 | 47.73 | 40 | 45.8 | 41 | 40.04 | 45 | 43 | 40.2 | 45 |  | 42.5537
| 1 | 5 | 37.4355 | 37.4 | 44 | 42 | 45 | 45 | 41.2686 | 48 | 42 | 46 | 40.8781 | 43 | 45.08 | 40 | 38.3 | 41 | 41.23 | 43 | 44 | 43 | 46 |  | 79
| 1 | 6 | 37.341 | 39.1 | 44 | 44 | 46 | 44 | 42.576 | 48 | 43 | 46 | 41.4212 | 42 | 44.99 | 40 | 38.4 | 41 | 40.93 | 46 | 45 | 42.6 | 44
| Mean |  | 37.983 | 39.0167 | 43.6667 | 42 | 45.1667 | 43.5 | 42.2057 | 49.5 | 43.1667 | 46 | 41.3247 | 42.6667 | 46.0967 | 40.1667 | 39.5833 | 40.6667 | 40.8333 | 44.5 | 44.3333 | 41.9667 | 45.1667
| Median |  | 37.9665 | 39.05 | 44 | 42 | 45 | 43.5 | 42.2052 | 49 | 43 | 46 | 41.4074 | 42.5 | 45.98 | 40 | 39.6 | 41 | 40.955 | 44.5 | 44.5 | 42.2 | 45.5
| Std Dev. |  | 0.57189 | 1.1321 | 0.5164 | 1.0954 | 0.75277 | 1.0488 | 0.70166 | 1.6432 | 1.472 | 0.63246 | 0.30078 | 0.8165 | 1.0558 | 0.40825 | 4.1835 | 0.8165 | 0.41428 | 1.0488 | 0.8165 | 0.9913 | 0.98319
| Rel.Std.Dev. |  | 1.506% | 2.902% | 1.183% | 2.608% | 1.667% | 2.411% | 1.662% | 3.32% | 3.41% | 1.375% | 0.7279% | 1.914% | 2.29% | 1.016% | 10.57% | 2.008% | 1.015% | 2.357% | 1.842% | 2.362% | 2.177%
| PDM3 |  | -10.74% | -8.312% | 2.616% | -1.301% | 6.14% | 2.224% | -0.8176% | 16.32% | 1.441% | 8.099% | -2.888% | 0.2655% | 8.326% | -5.609% | -6.98% | -4.434% | -4.043% | 4.574% | 4.182% | -1.379% | 6.14%
| Z-Score (Absolute) |  | 1.8 | 1.39 | 0.41 | 0.23 | 1 | 0.35 | 0.15 | 2.68 | 0 | 1.32 | 0.5 | 0.03 | 1.36 | 0.95 | 0.67 | 0.62 | 0.69 | 0.74 | 0.67 | 0.25 | 1
| Analytical results for Cr in OREAS 604 (Certified Value 32.1 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 21 | 23 | 24 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g | 40g |  |  |  | 2
| 1 | 1 | 31.9232 | 32 | 38 | 31 | 33 | 58 | 32.3569 | 31 | 32 | 37 | 34.2 | 30 | 29 | 30 | 28 | 35 | 33 | 30 | 25 |  |  |  | 1
| 1 | 2 | 32.637 | 32 | 29 | 33 | 37 | 33 | 32.7411 | 40 | 31 | 35 | 29 | 35 | 28 | 27 | 31 | 33 | 33 | 30 | 25 |  |  |  | 0
| 1 | 3 | 31.5878 | 31 | 33 | 31 | 32 | 29 | 33.0691 | 35 | 32 | 35 | 33.2 | 34 | 29 | 26 | 29 | 31 | 33 | 30 | 25 |  |  |  | 16
| 1 | 4 | 31.691 | 33 | 29 | 31 | 36 | 32 | 33.3661 | 31 | 30 | 35 | 33.3 | 36 | 29 | 27 | 32 | 31 | 29 | 30 | 25 |  |  |  | 32.0819
| 1 | 5 | 33.4712 | 32 | 33 | 31 | 35 | 32 | 35.0264 | 31 | 30 | 37 | 31.3 | 32 | 35 | 29 | 31 | 26 | 34 | 30 | 25 |  |  |  | 80
| 1 | 6 | 30.315 | 32 | 33 | 29 | 36 | 33 | 33.9604 | 40 | 30 | 40 | 32.1 | 34 | 27 | 28 | 30 | 29 | 34 | 40 | 25
| Mean |  | 31.9375 | 32 | 32.5 | 31 | 34.8333 | 36.1667 | 33.42 | 34.6667 | 30.8333 | 36.5 | 32.1833 | 33.5 | 29.5 | 27.8333 | 30.1667 | 30.8333 | 32.6667 | 31.6667 | 25
| Median |  | 31.8071 | 32 | 33 | 31 | 35.5 | 32.5 | 33.2176 | 33 | 30.5 | 36 | 32.65 | 34 | 29 | 27.5 | 30.5 | 31 | 33 | 30 | 25
| Std Dev. |  | 1.0637 | 0.63246 | 3.3317 | 1.2649 | 1.9408 | 10.7966 | 0.95807 | 4.4121 | 0.98319 | 1.9748 | 1.8584 | 2.1679 | 2.8107 | 1.472 | 1.472 | 3.1252 | 1.8619 | 4.0825 | 0
| Rel.Std.Dev. |  | 3.331% | 1.976% | 10.25% | 4.08% | 5.572% | 29.85% | 2.867% | 12.73% | 3.189% | 5.411% | 5.774% | 6.471% | 9.528% | 5.288% | 4.879% | 10.14% | 5.7% | 12.89% | 0%
| PDM3 |  | -0.45% | -0.2553% | 1.303% | -3.372% | 8.576% | 12.73% | 4.171% | 8.057% | -3.892% | 13.77% | 0.3162% | 4.42% | -8.048% | -13.24% | -5.97% | -3.892% | 1.823% | -1.294% | -22.07%
| Z-Score (Absolute) |  | 0 | 0.03 | 0.26 | 0.43 | 1.34 | 0.06 | 0.68 | 1.26 | 0.51 | 2.1 | 0.11 | 0.72 | 1.63 | 1.89 | 0.82 | 0.51 | 0.67 | 0.12 | 3.2
| Analytical results for Cs in OREAS 604 (Certified Value 1.22 ppm) | [23 empty] | Certified
Cs | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 10: 24; 11: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS |  |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 10: 40g; 11: 40g |  |  |  | 2
| 1 | 1 | 01: 1.45; 02: 1.502; 05: 1.4959; 06: 1.12; 07: 1.79; 08: 1.48; 09: 1.137; 10: 1; 11: 0.9 |  |  |  | 1
| 1 | 2 | 01: 1.39; 02: 1.5636; 05: 1.4004; 06: 1.17; 07: 1.8; 08: 1.2; 09: 1.152; 10: 1; 11: 0.96 |  |  |  | 0
| 1 | 3 | 01: 1.38; 02: 1.4976; 05: 1.4589; 06: 1.14; 07: 1.84; 08: 1.15; 09: 1.189; 10: 0.9; 11: 0.92 |  |  |  | 16
| 1 | 4 | 01: 1.41; 02: 1.6096; 05: 1.3943; 06: 1.18; 07: 1.92; 08: 1.14; 09: 1.124; 10: 0.9; 11: 0.92 |  |  |  | 1.2226
| 1 | 5 | 01: 1.36; 02: 1.5431; 05: 1.4992; 06: 1.15; 07: 1.71; 08: 1.3; 09: 1.142; 10: 0.9; 11: 0.96 |  |  |  | 81
| 1 | 6 | 01: 1.44; 02: 1.6071; 05: 1.4959; 06: 1.21; 07: 1.79; 08: 1.08; 09: 1.111; 10: 1; 11: 0.96
| Mean |  | 01: 1.405; 02: 1.5538; 05: 1.4574; 06: 1.1617; 07: 1.8083; 08: 1.225; 09: 1.1425; 10: 0.95; 11: 0.93667
| Median |  | 01: 1.4; 02: 1.5534; 05: 1.4774; 06: 1.16; 07: 1.795; 08: 1.175; 09: 1.1395; 10: 0.95; 11: 0.94
| Std Dev. |  | 01: 0.035071; 02: 0.049012; 05: 0.048882; 06: 0.031885; 07: 0.069113; 08: 0.14502; 09: 0.026898; 10: 0.054772; 11: 0.026583
| Rel.Std.Dev. |  | 01: 2.496%; 02: 3.154%; 05: 3.354%; 06: 2.745%; 07: 3.822%; 08: 11.84%; 09: 2.354%; 10: 5.766%; 11: 2.838%
| PDM3 |  | 01: 14.92%; 02: 27.09%; 05: 19.2%; 06: -4.987%; 07: 47.9%; 08: 0.1932%; 09: -6.555%; 10: -22.3%; 11: -23.39%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.11; 05: 0.83; 06: 0.04; 07: 1.85; 08: 0; 09: 0.09; 10: 0.65; 11: 0.69
| Analytical results for Cu in OREAS 604 (Certified Value 2.16 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | wt.%
|  |  | AR*OES | AR*OES | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*OES | AR*AAS | AR*AAS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | 2
|  |  | 0.5g | 40g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 20g | 01g | 0.5g | 01g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.5g | 0.25g | 01g | 40g | 40g | 3
| 1 | 1 | 2.1075 | > 10000 | 2.217 | 2.17 | 2.25 | 2.21 | 2.19 | 2.1162 | 2.09 | 2.2 | 2.0553 | 2.0639 | > 10000 | 2.22 | > 10000 | > 10000 | > 10000 | 2.1665 | 2.0513 | 2.189 | N.A. | 2.17 | 1
| 1 | 2 | 2.1044 | > 10000 | 2.159 | 2.18 | 2.24 | 2.18 | 2.2 | 2.0982 | 2.15 | 2.222 | 2.0791 | 2.0974 | > 10000 | 2.2 | > 10000 | > 10000 | > 10000 | 2.212 | 2.1163 | 2.127 | N.A. | 2.22 | 0
| 1 | 3 | 2.1051 | > 10000 | 2.141 | 2.13 | 2.32 | 2.06 | 2.24 | 2.1152 | 2.16 | 2.219 | 2.0565 | 2.0418 | > 10000 | 2.207 | > 10000 | > 10000 | > 10000 | 2.1934 | 2.1148 | 2.305 | N.A. | 2.21 | 16
| 1 | 4 | 2.1064 | > 10000 | 2.203 | 2.14 | 2.15 | 2.03 | 2.19 | 2.1248 | 2.08 | 2.151 | 2.1854 | 2.0913 | > 10000 | 2.257 | > 10000 | > 10000 | > 10000 | 2.1612 | 2.1234 | 2.125 | N.A. | 2.18 | 2.1579
| 1 | 5 | 2.0997 | > 10000 | 2.157 | 2.14 | 2.15 | 2.07 | 2.22 | 2.111 | 2.16 | 2.226 | 2.2315 | 2.0693 | > 10000 | 2.182 | > 10000 | > 10000 | -- | 2.1899 | 2.1218 | 2.195 | N.A. | 2.22 | 82
| 1 | 6 | 2.104 | > 10000 | 2.172 | 2.13 | 2.29 | 2.04 | 2.25 | 2.1338 | 2.14 | 2.289 | 2.0877 | 2.0099 | > 10000 | 2.165 | > 10000 | > 10000 | > 10000 | 2.1894 | 2.1578 | 2.196 | N.A. | 2.22
| Mean |  | 2.1045 |  | 2.1748 | 2.1483 | 2.2333 | 2.0983 | 2.215 | 2.1165 | 2.13 | 2.2178 | 2.1159 | 2.0623 |  | 2.2052 |  |  |  | 2.1854 | 2.1142 | 2.1895 |  | 2.2033
| Median |  | 2.1047 |  | 2.1655 | 2.14 | 2.245 | 2.065 | 2.21 | 2.1157 | 2.145 | 2.2205 | 2.0834 | 2.0666 |  | 2.2035 |  |  |  | 2.1896 | 2.119 | 2.192 |  | 2.215
| Std Dev. |  | 0.0027057 |  | 0.029301 | 0.02137 | 0.070616 | 0.07679 | 0.025884 | 0.012112 | 0.035777 | 0.044575 | 0.074221 | 0.032539 |  | 0.031896 |  |  |  | 0.018712 | 0.034671 | 0.065543 |  | 0.022509
| Rel.Std.Dev. |  | 0.1286% |  | 1.347% | 0.9947% | 3.162% | 3.66% | 1.169% | 0.5722% | 1.68% | 2.01% | 3.508% | 1.578% |  | 1.446% |  |  |  | 0.8562% | 1.64% | 2.994% |  | 1.022%
| PDM3 |  | -2.477% |  | 0.7827% | -0.4453% | 3.494% | -2.762% | 2.644% | -1.919% | -1.295% | 2.775% | -1.948% | -4.433% |  | 2.188% |  |  |  | 1.272% | -2.026% | 1.462% |  | 2.103%
| Z-Score (Absolute) |  | 0.87 | NA | 0.2 | 0.2 | 1.09 | 0.96 | 0.81 | 0.69 | 0.48 | 0.86 | 0.69 | 1.51 | NA | 0.66 | NA | NA | NA | 0.36 | 0.53 | 0.42 | NA | 0.64
| Analytical results for Dy in OREAS 604 (Certified Value 0.86 ppm) | [23 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 02; 02: 09; 04: 13; 05: 14; 06: 15; 07: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS | 2
|  |  | 01: 40g; 02: 20g; 04: 01g; 05: 0.5g; 06: 01g; 07: 40g | 2
| 1 | 1 | 01: 0.95; 02: 0.9519; 04: 0.88228; 05: 0.8; 06: 0.86; 07: 0.66 | 1
| 1 | 2 | 01: 0.91; 02: 0.96934; 04: 0.82631; 05: 0.8; 06: 0.96; 07: 0.66 | 0
| 1 | 3 | 01: 0.92; 02: 0.95634; 04: 0.84433; 05: 0.79; 06: 0.99; 07: 0.68 | 16
| 1 | 4 | 01: 0.92; 02: 0.96258; 04: 0.81505; 05: 0.8; 06: 1; 07: 0.68 | 0.86204
| 1 | 5 | 01: 0.93; 02: 0.96216; 04: 0.84564; 05: 0.82; 06: 1.01; 07: 0.66 | 83
| 1 | 6 | 01: 0.95; 02: 0.99628; 04: 0.83709; 05: 0.82; 06: 0.95; 07: 0.7
| Mean |  | 01: 0.93; 02: 0.96643; 04: 0.84178; 05: 0.805; 06: 0.96167; 07: 0.67333
| Median |  | 01: 0.925; 02: 0.96237; 04: 0.84071; 05: 0.8; 06: 0.975; 07: 0.67
| Std Dev. |  | 01: 0.016733; 02: 0.015781; 04: 0.022966; 05: 0.012247; 06: 0.054924; 07: 0.01633
| Rel.Std.Dev. |  | 01: 1.799%; 02: 1.633%; 04: 2.728%; 05: 1.521%; 06: 5.711%; 07: 2.425%
| PDM3 |  | 01: 7.883%; 02: 12.11%; 04: -2.35%; 05: -6.617%; 06: 11.56%; 07: -21.89%
| Z-Score (Absolute) |  | 01: 0.4; 02: 0.67; 04: 0.4; 05: 0.73; 06: 0.68; 07: 1.91
| Analytical results for Er in OREAS 604 (Certified Value 0.29 ppm) | [23 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 09; 04: 14; 05: 15; 06: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 40g; 02: 20g; 04: 0.5g; 05: 01g; 06: 40g | 2
| 1 | 1 | 01: 0.31; 02: 0.29875; 04: 0.28; 05: 0.32; 06: 0.24 | 1
| 1 | 2 | 01: 0.3; 02: 0.31255; 04: 0.28; 05: 0.36; 06: 0.24 | 0
| 1 | 3 | 01: 0.31; 02: 0.31571; 04: 0.27; 05: 0.29; 06: 0.24 | 16
| 1 | 4 | 01: 0.31; 02: 0.32171; 04: 0.26; 05: 0.31; 06: 0.24 | 0.28977
| 1 | 5 | 01: 0.31; 02: 0.30838; 04: 0.27; 05: 0.32; 06: 0.23 | 84
| 1 | 6 | 01: 0.31; 02: 0.30609; 04: 0.27; 05: 0.32; 06: 0.24
| Mean |  | 01: 0.30833; 02: 0.31053; 04: 0.27167; 05: 0.32; 06: 0.23833
| Median |  | 01: 0.31; 02: 0.31047; 04: 0.27; 05: 0.32; 06: 0.24
| Std Dev. |  | 01: 0.0040825; 02: 0.0079895; 04: 0.0075277; 05: 0.022804; 06: 0.0040825
| Rel.Std.Dev. |  | 01: 1.324%; 02: 2.573%; 04: 2.771%; 05: 7.126%; 06: 1.713%
| PDM3 |  | 01: 6.405%; 02: 7.164%; 04: -6.249%; 05: 10.43%; 06: -17.75%
| Z-Score (Absolute) |  | 01: 0; 02: 0.13; 04: 2.12; 05: 0.67; 06: 4.05
| Analytical results for Eu in OREAS 604 (Certified Value 0.39 ppm) | [23 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 09; 04: 14; 05: 15; 06: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 40g; 02: 20g; 04: 0.5g; 05: 01g; 06: 40g | 2
| 1 | 1 | 01: 0.41; 02: 0.41582; 04: 0.34; 05: 0.47; 06: 0.304 | 1
| 1 | 2 | 01: 0.39; 02: 0.43426; 04: 0.35; 05: 0.47; 06: 0.322 | 0
| 1 | 3 | 01: 0.4; 02: 0.42741; 04: 0.34; 05: 0.49; 06: 0.304 | 16
| 1 | 4 | 01: 0.39; 02: 0.42504; 04: 0.34; 05: 0.44; 06: 0.312 | 0.38892
| 1 | 5 | 01: 0.41; 02: 0.41089; 04: 0.35; 05: 0.5; 06: 0.312 | 85
| 1 | 6 | 01: 0.41; 02: 0.41431; 04: 0.34; 05: 0.43; 06: 0.316
| Mean |  | 01: 0.40167; 02: 0.42129; 04: 0.34333; 05: 0.46667; 06: 0.31167
| Median |  | 01: 0.405; 02: 0.42043; 04: 0.34; 05: 0.47; 06: 0.312
| Std Dev. |  | 01: 0.0098319; 02: 0.0090162; 04: 0.005164; 05: 0.027325; 06: 0.0069761
| Rel.Std.Dev. |  | 01: 2.448%; 02: 2.14%; 04: 1.504%; 05: 5.855%; 06: 2.238%
| PDM3 |  | 01: 3.276%; 02: 8.321%; 04: -11.72%; 05: 19.99%; 06: -19.86%
| Z-Score (Absolute) |  | 01: 0; 02: 0.23; 04: 0.67; 05: 0.75; 06: 1.04
| Analytical results for Fe in OREAS 604 (Certified Value 2.93 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 2.882 | 2.97 | 2.77 | 2.77 | 2.85 | 2.74 | 3.0855 | 2.93 | 2.89 | 3.03 | 3.0684 | 2.93 | 3.11 | 2.73 | 3.05 | 2.71 | 3.099 | 2.64 | 2.86 | 2.97 | 2.93 |  | 1
| 1 | 2 | 2.858 | 2.95 | 2.81 | 2.71 | 2.84 | 2.78 | 3.182 | 3.11 | 3.08 | 3.07 | 2.9903 | 3 | 3.13 | 2.79 | 2.83 | 2.76 | 3.204 | 2.7 | 2.85 | 3.07 | 2.95 |  | 0
| 1 | 3 | 2.87 | 2.98 | 2.82 | 2.72 | 2.82 | 2.78 | 3.0941 | 3.13 | 3.2 | 3.06 | 2.9659 | 2.96 | 3.13 | 2.79 | 2.96 | 2.71 | 3.25 | 2.7 | 2.86 | 2.99 | 2.98 |  | 16
| 1 | 4 | 2.848 | 2.92 | 2.83 | 2.72 | 2.83 | 2.79 | 3.1762 | 2.97 | 3.05 | 3.02 | 2.9875 | 2.9 | 3.21 | 2.84 | 2.98 | 2.71 | 3.176 | 2.72 | 2.78 | 3.26 | 2.95 |  | 2.9267
| 1 | 5 | 2.792 | 2.92 | 2.81 | 2.73 | 2.81 | 2.85 | 3.0621 | 3.02 | 3.06 | 2.97 | 3.0149 | 2.98 | 3.07 | 2.8 | 2.98 | 2.76 | 3.216 | 2.7 | 2.81 | 2.92 | 2.99 |  | 86
| 1 | 6 | 2.814 | 2.91 | 2.81 | 2.89 | 2.83 | 2.81 | 3.1133 | 3.04 | 2.96 | 3.1 | 3.007 | 2.93 | 3.07 | 2.81 | 2.91 | 2.64 | 3.184 | 2.74 | 2.91 | 3.12 | 2.95
| Mean |  | 2.844 | 2.9417 | 2.8083 | 2.7567 | 2.83 | 2.7917 | 3.1189 | 3.0333 | 3.04 | 3.0417 | 3.0057 | 2.95 | 3.12 | 2.7933 | 2.9517 | 2.715 | 3.1882 | 2.7 | 2.845 | 3.055 | 2.9583
| Median |  | 2.853 | 2.935 | 2.81 | 2.725 | 2.83 | 2.785 | 3.1037 | 3.03 | 3.055 | 3.045 | 2.9986 | 2.945 | 3.12 | 2.795 | 2.97 | 2.71 | 3.194 | 2.7 | 2.855 | 3.03 | 2.95
| Std Dev. |  | 0.034456 | 0.029269 | 0.020412 | 0.068605 | 0.014142 | 0.03656 | 0.049498 | 0.077632 | 0.1064 | 0.04535 | 0.035107 | 0.036878 | 0.051769 | 0.036148 | 0.074677 | 0.044159 | 0.050921 | 0.033466 | 0.045056 | 0.12341 | 0.022286
| Rel.Std.Dev. |  | 1.212% | 0.995% | 0.7269% | 2.489% | 0.4997% | 1.31% | 1.587% | 2.559% | 3.5% | 1.491% | 1.168% | 1.25% | 1.659% | 1.294% | 2.53% | 1.626% | 1.597% | 1.239% | 1.584% | 4.04% | 0.7533%
| PDM3 |  | -2.827% | 0.5097% | -4.046% | -5.811% | -3.306% | -4.615% | 6.565% | 3.642% | 3.87% | 3.926% | 2.696% | 0.7944% | 6.603% | -4.558% | 0.8514% | -7.235% | 8.932% | -7.747% | -2.793% | 4.382% | 1.079%
| Z-Score (Absolute) |  | 0.67 | 0.05 | 0.9 | 1.4 | 0.76 | 1.01 | 1.07 | 0.53 | 0.57 | 0.58 | 0.35 | 0 | 1.08 | 1 | 0.01 | 1.49 | 1.52 | 1.59 | 0.67 | 0.67 | 0.05
| Analytical results for Ga in OREAS 604 (Certified Value 6.1 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 17 | 19 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 40g | 40g | [6 empty] | 2
| 1 | 1 | 6.6888 | 5.79 | < 10 | 10 | 10 | 10 | 7.121 | 5 | 5.908 | 5.21 | 6 | 6.9 | 4.82 | < 5 | 6.2 | 5.6 | [6 empty] | 1
| 1 | 2 | 6.6912 | 5.94 | < 10 | < 10 | 10 | 10 | 7.1938 | 5 | 6.014 | 5.13 | 5 | 5.6 | 4.97 | 8 | 6.1 | 5.8 | [6 empty] | 0
| 1 | 3 | 6.7776 | 5.8 | < 10 | 10 | 10 | < 10 | 7.1052 | 6 | 6.0408 | 5.16 | 6 | 6.3 | 5.04 | < 5 | 5.6 | 5.6 | [6 empty] | 16
| 1 | 4 | 6.456 | 5.78 | < 10 | 10 | 10 | 10 | 6.9085 | 8 | 5.9334 | 5.11 | 5 | 6.3 | 4.85 | < 5 | 5.8 | 5.6 | [6 empty] | 6.0957
| 1 | 5 | 6.66 | 5.79 | < 10 | < 10 | 10 | 10 | 6.788 | 7 | 5.7031 | 5.09 | 4 | 6.3 | 4.96 | < 5 | 6.1 | 5.8 | [6 empty] | 87
| 1 | 6 | 6.3264 | 5.91 | < 10 | 10 | 10 | 10 | 6.8615 | 8 | 6.0139 | 5.09 | 6 | 5.6 | 4.98 | 8 | 6 | 5.8
| Mean |  | 6.6 | 5.835 |  | 10 | 10 | 10 | 6.9963 | 6.5 | 5.9355 | 5.1317 | 5.3333 | 6.1667 | 4.9367 | 8 | 5.9667 | 5.7
| Median |  | 6.6744 | 5.795 |  | 10 | 10 | 10 | 7.0069 | 6.5 | 5.9736 | 5.12 | 5.5 | 6.3 | 4.965 | 8 | 6.05 | 5.7
| Std Dev. |  | 0.17143 | 0.07064 |  | 0 | 0 | 0 | 0.16474 | 1.3784 | 0.125 | 0.046655 | 0.8165 | 0.49666 | 0.084063 | 0 | 0.22509 | 0.10954
| Rel.Std.Dev. |  | 2.597% | 1.211% |  | 0% | 0% | 0% | 2.355% | 21.21% | 2.106% | 0.9092% | 15.31% | 8.054% | 1.703% | 0% | 3.773% | 1.922%
| PDM3 |  | 8.273% | -4.277% |  | 64.05% | 64.05% | 64.05% | 14.78% | 6.633% | -2.627% | -15.81% | -12.51% | 1.164% | -19.01% | 31.24% | -2.117% | -6.491%
| Z-Score (Absolute) |  | 0.35 | 0.27 | NA | 3.12 | 3.12 | 3.12 | 0.67 | 0.27 | 0.15 | 0.84 | 0.68 | 0 | 1 | 1.49 | 0.16 | 0.38
| Analytical results for Gd in OREAS 604 (Certified Value 1.49 ppm) | [23 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  | 2
| 1 | 1 | 01: 1.58; 02: 1.5779; 05: 1.3497; 06: 1.42; 07: 1.82; 08: 1.13 |  | 1
| 1 | 2 | 01: 1.53; 02: 1.5735; 05: 1.3646; 06: 1.4; 07: 1.84; 08: 1.19 |  | 0
| 1 | 3 | 01: 1.53; 02: 1.5849; 05: 1.4029; 06: 1.42; 07: 1.8; 08: 1.17 |  | 16
| 1 | 4 | 01: 1.55; 02: 1.5283; 05: 1.3776; 06: 1.37; 07: 1.87; 08: 1.13 |  | 1.4851
| 1 | 5 | 01: 1.59; 02: 1.5386; 05: 1.4399; 06: 1.4; 07: 1.71; 08: 1.16 |  | 88
| 1 | 6 | 01: 1.59; 02: 1.6064; 05: 1.388; 06: 1.43; 07: 1.9; 08: 1.2
| Mean |  | 01: 1.5617; 02: 1.5683; 05: 1.3871; 06: 1.4067; 07: 1.8233; 08: 1.1633
| Median |  | 01: 1.565; 02: 1.5757; 05: 1.3828; 06: 1.41; 07: 1.83; 08: 1.165
| Std Dev. |  | 01: 0.028577; 02: 0.029427; 05: 0.031707; 06: 0.021602; 07: 0.065929; 08: 0.029439
| Rel.Std.Dev. |  | 01: 1.83%; 02: 1.876%; 05: 2.286%; 06: 1.536%; 07: 3.616%; 08: 2.531%
| PDM3 |  | 01: 5.158%; 02: 5.603%; 05: -6.596%; 06: -5.279%; 07: 22.78%; 08: -21.66%
| Z-Score (Absolute) |  | 01: 0.58; 02: 0.63; 05: 0.72; 06: 0.58; 07: 2.52; 08: 2.39
| Analytical results for Ge in OREAS 604 (Indicative Value 0.2 ppm) | [23 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 11; 02: 15; 05: 17; 06: 19 |  | ppm
|  |  | 01: AR*OES; 02: AR*OES/MS; 05: AR*OES/MS; 06: AR*OES/MS |  | 2
|  |  | 01: 0.15g; 02: 01g; 05: 0.25g; 06: 0.25g |  | 2
| 1 | 1 | 01: < 10; 02: < 0.5; 05: 0.3; 06: 0.202 |  | 1
| 1 | 2 | 01: < 10; 02: < 0.5; 05: 0.2; 06: 0.198 |  | 7
| 1 | 3 | 01: < 10; 02: < 0.5; 05: 0.2; 06: 0.241 |  | 16
| 1 | 4 | 01: < 10; 02: < 0.5; 05: 0.1; 06: 0.19 |  | 0.2035
| 1 | 5 | 01: < 10; 02: < 0.5; 05: 0.2; 06: 0.22 |  | 13
| 1 | 6 | 01: < 10; 02: < 0.5; 05: 0.2; 06: 0.191
| Mean |  | 05: 0.2; 06: 0.207
| Median |  | 05: 0.2; 06: 0.2
| Std Dev. |  | 05: 0.063246; 06: 0.01988
| Rel.Std.Dev. |  | 05: 31.62%; 06: 9.604%
| PDM3 |  | 05: -1.72%; 06: 1.72%
| Z-Score (Absolute) |  | 01: NA; 02: NA; 05: 0.67; 06: 0.67
| Analytical results for Hf in OREAS 604 (Certified Value 0.58 ppm) | [23 empty] | Certified
Hf | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 10: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 10: 40g |  | 2
| 1 | 1 | 01: 0.75; 02: 0.64587; 05: 0.42686; 06: 0.53; 07: 0.77; 08: 0.51; 09: 0.527; 10: 0.4 |  | 1
| 1 | 2 | 01: 0.73; 02: 0.63648; 05: 0.41608; 06: 0.53; 07: 0.72; 08: 0.45; 09: 0.53; 10: 0.6 |  | 0
| 1 | 3 | 01: 0.73; 02: 0.67148; 05: 0.44253; 06: 0.53; 07: 0.59; 08: 0.56; 09: 0.548; 10: 0.4 |  | 16
| 1 | 4 | 01: 0.73; 02: 0.71526; 05: 0.42353; 06: 0.53; 07: 0.74; 08: 0.44; 09: 0.515; 10: 0.6 |  | 0.57778
| 1 | 5 | 01: 0.71; 02: 0.64375; 05: 0.42044; 06: 0.54; 07: 0.8; 08: 0.49; 09: 0.523; 10: 0.6 |  | 89
| 1 | 6 | 01: 0.76; 02: 0.68436; 05: 0.44896; 06: 0.55; 07: 0.61; 08: 0.48; 09: 0.535; 10: 0.6
| Mean |  | 01: 0.735; 02: 0.6662; 05: 0.42973; 06: 0.535; 07: 0.705; 08: 0.48833; 09: 0.52967; 10: 0.53333
| Median |  | 01: 0.73; 02: 0.65867; 05: 0.42519; 06: 0.53; 07: 0.73; 08: 0.485; 09: 0.5285; 10: 0.6
| Std Dev. |  | 01: 0.017607; 02: 0.030214; 05: 0.013059; 06: 0.0083666; 07: 0.085965; 08: 0.043551; 09: 0.011237; 10: 0.10328
| Rel.Std.Dev. |  | 01: 2.395%; 02: 4.535%; 05: 3.039%; 06: 1.564%; 07: 12.19%; 08: 8.918%; 09: 2.121%; 10: 19.36%
| PDM3 |  | 01: 27.21%; 02: 15.3%; 05: -25.62%; 06: -7.405%; 07: 22.02%; 08: -15.48%; 09: -8.328%; 10: -7.693%
| Z-Score (Absolute) |  | 01: 1.8; 02: 1.18; 05: 0.94; 06: 0.01; 07: 1.53; 08: 0.41; 09: 0.04; 10: 0.01
| Analytical results for Hg in OREAS 604 (Certified Value 1.01 ppm) | [23 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 16 | 17 | 19 | 23 | 24 | 25 | [8 empty] | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | [8 empty] | 2
|  |  | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g | [8 empty] | 3
| 1 | 1 | 1 | 1 | 2 | 1 | 0.95205 | 1 | < 3 | 1.05 | < 5 | 2.05 | 0.881 | 1 | 1 | 1.06 | [8 empty] | 1
| 1 | 2 | 1 | 1 | 1 | 1 | 1.0035 | 1 | < 3 | 1.02 | < 5 | 1.69 | 0.955 | 1 | 1.5 | 1.15 | [8 empty] | 3
| 1 | 3 | 1 | 1 | 1 | 1 | 0.98143 | 1 | < 3 | 1.05 | < 5 | 1.79 | 0.965 | 1 | 1 | 1.14 | [8 empty] | 16
| 1 | 4 | 1 | 2 | 1 | 1 | 0.98467 | 1 | < 3 | 1.03 | < 5 | 2.11 | 1.01 | 1 | 1 | 1.08 | [8 empty] | 1.0073
| 1 | 5 | 1 | 2 | 2 | 1 | 0.9362 | 1 | < 3 | 1.06 | < 5 | 1.84 | 0.876 | 1 | 1 | 1.15 | [8 empty] | 90
| 1 | 6 | 1 | 1 | 1 | < 1 | 0.99879 | 1 | < 3 | 1.02 | < 5 | 1.7 | 0.949 | < 1 | 1 | 1.18
| Mean |  | 01: 1; 02: 1.3333; 05: 1.3333; 06: 1; 07: 0.97611; 08: 1; 10: 1.0383; 12: 1.8633; 13: 0.93933; 14: 1; 15: 1.0833; 16: 1.1267
| Median |  | 01: 1; 02: 1; 05: 1; 06: 1; 07: 0.98305; 08: 1; 10: 1.04; 12: 1.815; 13: 0.952; 14: 1; 15: 1; 16: 1.145
| Std Dev. |  | 01: 0; 02: 0.5164; 05: 0.5164; 06: 0; 07: 0.026604; 08: 0; 10: 0.017224; 12: 0.17795; 13: 0.051779; 14: 0; 15: 0.20412; 16: 0.046332
| Rel.Std.Dev. |  | 01: 0%; 02: 38.73%; 05: 38.73%; 06: 0%; 07: 2.726%; 08: 0%; 10: 1.659%; 12: 9.55%; 13: 5.512%; 14: 0%; 15: 18.84%; 16: 4.112%
| PDM3 |  | 01: -0.726%; 02: 32.37%; 05: 32.37%; 06: -0.726%; 07: -3.098%; 08: -0.726%; 10: 3.08%; 12: 84.98%; 13: -6.749%; 14: -0.726%; 15: 7.547%; 16: 11.85%
Z-Score (Absolute) | bar: 01 to 16
| Analytical results for Ho in OREAS 604 (Certified Value 0.14 ppm) | [23 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  | 3
| 1 | 1 | 01: 0.14; 02: 0.14526; 05: 0.14755; 06: 0.13; 07: 0.15; 08: 0.105 |  | 1
| 1 | 2 | 01: 0.14; 02: 0.13946; 05: 0.13625; 06: 0.13; 07: 0.14; 08: 0.11 |  | 0
| 1 | 3 | 01: 0.14; 02: 0.14939; 05: 0.14054; 06: 0.13; 07: 0.14; 08: 0.11 |  | 16
| 1 | 4 | 01: 0.14; 02: 0.14358; 05: 0.14115; 06: 0.13; 07: 0.14; 08: 0.11 |  | 0.1393
| 1 | 5 | 01: 0.14; 02: 0.13538; 05: 0.14725; 06: 0.13; 07: 0.14; 08: 0.11 |  | 91
| 1 | 6 | 01: 0.14; 02: 0.14889; 05: 0.14445; 06: 0.13; 07: 0.13; 08: 0.11
| Mean |  | 01: 0.14; 02: 0.14366; 05: 0.14286; 06: 0.13; 07: 0.14; 08: 0.10917
| Median |  | 01: 0.14; 02: 0.14442; 05: 0.1428; 06: 0.13; 07: 0.14; 08: 0.11
| Std Dev. |  | 01: 0; 02: 0.0054578; 05: 0.0043759; 06: 0; 07: 0.0063246; 08: 0.0020412
| Rel.Std.Dev. |  | 01: 0%; 02: 3.799%; 05: 3.063%; 06: 0%; 07: 4.518%; 08: 1.87%
| PDM3 |  | 01: 0.4992%; 02: 3.127%; 05: 2.554%; 06: -6.679%; 07: 0.4992%; 08: -21.63%
| Z-Score (Absolute) |  | 01: 0; 02: 0.76; 05: 0.59; 06: 2.07; 07: 0; 08: 6.38
| Analytical results for In in OREAS 604 (Certified Value 3.66 ppm) | [23 empty] | Certified
In | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 13; 08: 14; 09: 15; 10: 17; 11: 19; 12: 24; 13: 25 |  | ppm
|  |  | 01: AR*OES; 02: AR*MS; 05: AR*OES/MS; 06: AR*OES; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*OES/MS; 11: AR*OES/MS; 12: AR*MS; 13: AR*MS |  | 2
|  |  | 01: 0.5g; 02: 40g; 05: 20g; 06: 0.15g; 07: 01g; 08: 0.5g; 09: 01g; 10: 0.25g; 11: 0.25g; 12: 40g; 13: 40g |  | 2
| 1 | 1 | 01: < 5; 02: 3.53; 05: 3.7225; 06: < 10; 07: 4.1881; 08: 3.57; 09: 4.28; 10: 3.5; 11: 3.239; 12: 4; 13: 3.31 |  | 1
| 1 | 2 | 01: < 5; 02: 3.54; 05: 3.7699; 06: < 10; 07: 4.163; 08: 3.72; 09: 4.34; 10: 2.84; 11: 3.202; 12: 4.5; 13: 3.37 |  | 0
| 1 | 3 | 01: < 5; 02: 3.48; 05: 3.7195; 06: < 10; 07: 4.1981; 08: 3.63; 09: 4.27; 10: 3.11; 11: 3.2; 12: 4; 13: 3.33 |  | 16
| 1 | 4 | 01: < 5; 02: 3.45; 05: 3.6446; 06: < 10; 07: 4.1866; 08: 3.53; 09: 4.31; 10: 3.66; 11: 3.061; 12: 4; 13: 3.32 |  | 3.6626
| 1 | 5 | 01: < 5; 02: 3.48; 05: 3.623; 06: < 10; 07: 4.2619; 08: 3.67; 09: 4.14; 10: 3.22; 11: 3.162; 12: 4; 13: 3.32 |  | 92
| 1 | 6 | 01: < 5; 02: 3.55; 05: 3.701; 06: < 10; 07: 4.1381; 08: 3.59; 09: 3.9; 10: 2.88; 11: 3.06; 12: 4; 13: 3.33
| Mean |  | 02: 3.505; 05: 3.6968; 07: 4.1893; 08: 3.6183; 09: 4.2067; 10: 3.2017; 11: 3.154; 12: 4.0833; 13: 3.33
| Median |  | 02: 3.505; 05: 3.7102; 07: 4.1873; 08: 3.61; 09: 4.275; 10: 3.165; 11: 3.181; 12: 4; 13: 3.325
| Std Dev. |  | 02: 0.040373; 05: 0.054262; 07: 0.041657; 08: 0.069402; 09: 0.16513; 10: 0.32927; 11: 0.076412; 12: 0.20412; 13: 0.020976
| Rel.Std.Dev. |  | 02: 1.152%; 05: 1.468%; 07: 0.9944%; 08: 1.918%; 09: 3.925%; 10: 10.28%; 11: 2.423%; 12: 4.999%; 13: 0.6299%
| PDM3 |  | 02: -4.302%; 05: 0.9335%; 07: 14.38%; 08: -1.208%; 09: 14.86%; 10: -12.58%; 11: -13.89%; 12: 11.49%; 13: -9.08%
| Z-Score (Absolute) |  | 01: NA; 02: 0.2; 05: 0.14; 06: NA; 07: 1.01; 08: 0; 09: 1.15; 10: 0.74; 11: 0.82; 12: 0.67; 13: 0.51
| Analytical results for K in OREAS 604 (Certified Value 0.167 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  | 3
| 1 | 1 | 0.25 | 0.2355 | 0.18 | 0.16 | 0.17 | 0.15 | 0.24719 | 0.18 | 0.2 | 0.16955 | 0.17 | 0.2 | 0.15 | 0.2 | 0.13 | 0.149 | 0.14 | 0.15 | 0.16 | 0.13 |  |  | 1
| 1 | 2 | 0.24 | 0.2307 | 0.17 | 0.15 | 0.17 | 0.16 | 0.25076 | 0.17 | 0.2 | 0.17659 | 0.18 | 0.22 | 0.16 | 0.18 | 0.14 | 0.157 | 0.14 | 0.15 | 0.18 | 0.13 |  |  | 0
| 1 | 3 | 0.24 | 0.2248 | 0.17 | 0.16 | 0.17 | 0.16 | 0.24237 | 0.17 | 0.2 | 0.1701 | 0.17 | 0.22 | 0.16 | 0.2 | 0.13 | 0.164 | 0.15 | 0.16 | 0.16 | 0.13 |  |  | 16
| 1 | 4 | 0.23 | 0.2417 | 0.17 | 0.15 | 0.17 | 0.16 | 0.24769 | 0.18 | 0.2 | 0.17864 | 0.17 | 0.22 | 0.16 | 0.19 | 0.14 | 0.156 | 0.15 | 0.15 | 0.17 | 0.13 |  |  | 0.16692
| 1 | 5 | 0.24 | 0.24 | 0.17 | 0.15 | 0.17 | 0.17 | 0.23965 | 0.18 | 0.19 | 0.17564 | 0.17 | 0.2 | 0.16 | 0.19 | 0.14 | 0.16 | 0.15 | 0.15 | 0.18 | 0.14 |  |  | 93
| 1 | 6 | 0.21 | 0.2392 | 0.17 | 0.16 | 0.17 | 0.16 | 0.24806 | 0.18 | 0.2 | 0.17903 | 0.17 | 0.21 | 0.16 | 0.19 | 0.13 | 0.16 | 0.15 | 0.16 | 0.18 | 0.14
| Mean |  | 0.235 | 0.23532 | 0.17167 | 0.155 | 0.17 | 0.16 | 0.24595 | 0.17667 | 0.19833 | 0.17492 | 0.17167 | 0.21167 | 0.15833 | 0.19167 | 0.135 | 0.15767 | 0.14667 | 0.15333 | 0.17167 | 0.13333
| Median |  | 0.24 | 0.23735 | 0.17 | 0.155 | 0.17 | 0.16 | 0.24744 | 0.18 | 0.2 | 0.17611 | 0.17 | 0.215 | 0.16 | 0.19 | 0.135 | 0.1585 | 0.15 | 0.15 | 0.175 | 0.13
| Std Dev. |  | 0.013784 | 0.0064762 | 0.0040825 | 0.0054772 | 0 | 0.0063246 | 0.0041126 | 0.005164 | 0.0040825 | 0.0041509 | 0.0040825 | 0.0098319 | 0.0040825 | 0.0075277 | 0.0054772 | 0.0050859 | 0.005164 | 0.005164 | 0.0098319 | 0.005164
| Rel.Std.Dev. |  | 5.866% | 2.752% | 2.378% | 3.534% | 0% | 3.953% | 1.672% | 2.923% | 2.058% | 2.373% | 2.378% | 4.645% | 2.578% | 3.928% | 4.057% | 3.226% | 3.521% | 3.368% | 5.727% | 3.873%
| PDM3 |  | 40.79% | 40.98% | 2.846% | -7.139% | 1.847% | -4.144% | 47.35% | 5.841% | 18.82% | 4.796% | 2.846% | 26.81% | -5.142% | 14.83% | -19.12% | -5.542% | -12.13% | -8.138% | 2.846% | -20.12%
| Z-Score (Absolute) |  | 2.44 | 2.45 | 0 | 0.64 | 0.06 | 0.45 | 2.86 | 0.19 | 1.03 | 0.13 | 0 | 1.54 | 0.51 | 0.77 | 1.41 | 0.54 | 0.96 | 0.71 | 0 | 1.48
| Analytical results for La in OREAS 604 (Certified Value 10.5 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 25 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g |  |  |  |  | 1
| 1 | 1 | 9.3168 | 12.89 | 10 | 10 | 10 | 10 | 13.6841 | 12 | 10 | 9.0825 | 10 | 14 | 9.7 | 10.6 | < 10 | 10.41 | 11 | 5.85 |  |  |  |  | 1
| 1 | 2 | 8.8632 | 12.3 | 10 | 10 | 10 | 10 | 13.956 | 12 | 11 | 8.7078 | 10 | 14 | 9.8 | 7.7 | < 10 | 10.51 | 11 | 6.32 |  |  |  |  | 0
| 1 | 3 | 8.928 | 12.22 | 10 | 10 | 10 | 10 | 13.4943 | 12 | 11 | 8.6966 | 10 | 14 | 10.6 | 9.8 | < 10 | 10.86 | 11 | 6.06 |  |  |  |  | 16
| 1 | 4 | 9.0576 | 12.43 | 10 | 10 | 10 | 10 | 13.8164 | 13 | 11 | 8.2644 | 10 | 15 | 9.9 | 11.1 | < 10 | 10.12 | 11 | 6.34 |  |  |  |  | 10.5317
| 1 | 5 | 8.982 | 12.38 | 10 | 10 | 10 | 10 | 13.0745 | 12 | 10 | 9.2082 | 10 | 14 | 9.9 | 10 | < 10 | 10.3 | 12 | 6.12 |  |  |  |  | 94
| 1 | 6 | 8.8128 | 12.56 | 10 | 10 | 10 | 10 | 13.7988 | 13 | 10 | 9.1367 | 10 | 14 | 10.4 | 11.1 | < 10 | 9.99 | 12 | 6.42
| Mean |  | 8.9934 | 12.4633 | 10 | 10 | 10 | 10 | 13.6374 | 12.3333 | 10.5 | 8.8494 | 10 | 14.1667 | 10.05 | 10.05 |  | 10.365 | 11.3333 | 6.185
| Median |  | 8.955 | 12.405 | 10 | 10 | 10 | 10 | 13.7415 | 12 | 10.5 | 8.8952 | 10 | 14 | 9.9 | 10.3 |  | 10.355 | 11 | 6.22
| Std Dev. |  | 0.18038 | 0.23889 | 0 | 0 | 0 | 0 | 0.31593 | 0.5164 | 0.54772 | 0.36092 | 0 | 0.40825 | 0.36194 | 1.2724 |  | 0.30755 | 0.5164 | 0.21427
| Rel.Std.Dev. |  | 2.006% | 1.917% | 0% | 0% | 0% | 0% | 2.317% | 4.187% | 5.216% | 4.079% | 0% | 2.882% | 3.601% | 12.66% |  | 2.967% | 4.556% | 3.464%
| PDM3 |  | -14.61% | 18.34% | -5.049% | -5.049% | -5.049% | -5.049% | 29.49% | 17.11% | -0.3013% | -15.97% | -5.049% | 34.51% | -4.574% | -4.574% |  | -1.583% | 7.611% | -41.27%
| Z-Score (Absolute) |  | 1.58 | 3.62 | 0.07 | 0.07 | 0.07 | 0.07 | 5.54 | 3.42 | 0.67 | 1.8 | 0.07 | 6.17 | 0 | 0 | NA | 0.47 | 1.92 | 5.79
| Analytical results for Li in OREAS 604 (Certified Value 6.16 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [9 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 40g | [9 empty] | 2
| 1 | 1 | 8.261 | 6.4 | 8.8573 | 6 | 5.1902 | 6.4 | 6 | 5 | 7 | 6 | 6.875 | 6 | 5.2 | [9 empty] | 1
| 1 | 2 | 7.887 | 6.5 | 8.9028 | 6 | 5.0054 | 6.2 | 5 | 6 | 7 | 7 | 7.161 | 6 | 5.4 | [9 empty] | 0
| 1 | 3 | 7.7605 | 6.2 | 8.88 | 5 | 5.0325 | 6.5 | 5 | 6 | 7 | 7 | 7.499 | 6 | 5.3 | [9 empty] | 16
| 1 | 4 | 7.9145 | 6.3 | 8.9392 | 5 | 5.4046 | 5.8 | 5 | 6 | 7 | 6 | 7.096 | 6 | 5.4 | [9 empty] | 6.1564
| 1 | 5 | 7.755 | 6.3 | 8.8602 | 5 | 5.3371 | 5.8 | 5 | 6 | 7 | 6 | 7.293 | 6 | 5.6 | [9 empty] | 95
| 1 | 6 | 7.6945 | 6.4 | 9.1425 | 5 | 5.2585 | 6.1 | 5 | 6 | 7 | 6 | 7.495 | 6 | 6
| Mean |  | 7.8788 | 6.35 | 8.9303 | 5.3333 | 5.2047 | 6.1333 | 5.1667 | 5.8333 | 7 | 6.3333 | 7.2365 | 6 | 5.4833
| Median |  | 7.8238 | 6.35 | 8.8914 | 5 | 5.2243 | 6.15 | 5 | 6 | 7 | 6 | 7.227 | 6 | 5.4
| Std Dev. |  | 0.20528 | 0.10488 | 0.10832 | 0.5164 | 0.16124 | 0.29439 | 0.40825 | 0.40825 | 0 | 0.5164 | 0.24292 | 0 | 0.28577
| Rel.Std.Dev. |  | 2.606% | 1.652% | 1.213% | 9.682% | 3.098% | 4.8% | 7.902% | 6.999% | 0% | 8.154% | 3.357% | 0% | 5.212%
| PDM3 |  | 27.98% | 3.145% | 45.06% | -13.37% | -15.46% | -0.3747% | -16.08% | -5.248% | 13.7% | 2.874% | 17.54% | -2.54% | -10.93%
| Z-Score (Absolute) |  | 1.41 | 0.18 | 2.36 | 0.67 | 0.78 | 0 | 0.81 | 0.25 | 0.73 | 0.17 | 0.93 | 0.11 | 0.55
| Analytical results for Lu in OREAS 604 (Certified Value 27 ppb) | [23 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 14; 06: 15; 07: 17; 08: 19; 09: 25 |  |  |  |  | ppb
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*MS |  |  |  |  | 0
|  |  | 01: 40g; 02: 20g; 05: 0.5g; 06: 01g; 07: 0.25g; 08: 0.25g; 09: 40g |  |  |  |  | 0
| 1 | 1 | 01: 30; 02: 26.7762; 05: 29; 06: < 100; 07: 30; 08: 29; 09: 20 |  |  |  |  | 1
| 1 | 2 | 01: 30; 02: 29.65; 05: 28; 06: < 100; 07: 30; 08: 28; 09: 20 |  |  |  |  | 0
| 1 | 3 | 01: 30; 02: 30.6537; 05: 28; 06: < 100; 07: 30; 08: 28; 09: 20 |  |  |  |  | 16
| 1 | 4 | 01: 30; 02: 27.8545; 05: 29; 06: < 100; 07: 20; 08: 27; 09: 20 |  |  |  |  | 27.3341
| 1 | 5 | 01: 40; 02: 25.7318; 05: 27; 06: < 100; 07: 30; 08: 29; 09: 20 |  |  |  |  | 96
| 1 | 6 | 01: 30; 02: 27.3629; 05: 26; 06: < 100; 07: 30; 08: 28; 09: 20
| Mean |  | 01: 31.6667; 02: 28.0048; 05: 27.8333; 07: 28.3333; 08: 28.1667; 09: 20
| Median |  | 01: 30; 02: 27.6087; 05: 28; 07: 30; 08: 28; 09: 20
| Std Dev. |  | 01: 4.0825; 02: 1.8349; 05: 1.169; 07: 4.0825; 08: 0.75277; 09: 0
| Rel.Std.Dev. |  | 01: 12.89%; 02: 6.552%; 05: 4.2%; 07: 14.41%; 08: 2.673%; 09: 0%
| PDM3 |  | 01: 15.85%; 02: 2.454%; 05: 1.826%; 07: 3.655%; 08: 3.046%; 09: -26.83%
| Z-Score (Absolute) |  | 01: 9.66; 02: 0.22; 05: 0.68; 06: NA; 07: 0.67; 08: 0.22; 09: 21.81
| Analytical results for Mg in OREAS 604 (Certified Value 0.106 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 0.13333 | 0.13 | 0.1 | 0.1 | 0.1 | 0.09 | 0.13118 | 0.08 | 0.1 | 0.12 | 0.10055 | 0.1 | 0.11 | 0.09 | 0.12 | 0.09 | 0.123 | 0.09 | 0.11 | 0.12 | 0.09 |  | 1
| 1 | 2 | 0.13333 | 0.12 | 0.1 | 0.09 | 0.1 | 0.1 | 0.1327 | 0.09 | 0.1 | 0.12 | 0.10061 | 0.1 | 0.11 | 0.1 | 0.11 | 0.09 | 0.125 | 0.09 | 0.11 | 0.13 | 0.09 |  | 0
| 1 | 3 | 0.13333 | 0.12 | 0.1 | 0.09 | 0.1 | 0.1 | 0.13001 | 0.09 | 0.1 | 0.12 | 0.1011 | 0.1 | 0.11 | 0.1 | 0.12 | 0.09 | 0.131 | 0.09 | 0.11 | 0.11 | 0.09 |  | 16
| 1 | 4 | 0.13333 | 0.13 | 0.1 | 0.1 | 0.1 | 0.1 | 0.13313 | 0.085 | 0.1 | 0.12 | 0.10078 | 0.1 | 0.11 | 0.1 | 0.11 | 0.09 | 0.124 | 0.09 | 0.1 | 0.12 | 0.09 |  | 0.10638
| 1 | 5 | 0.13333 | 0.13 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1282 | 0.085 | 0.1 | 0.12 | 0.10189 | 0.1 | 0.1 | 0.09 | 0.11 | 0.09 | 0.129 | 0.09 | 0.11 | 0.11 | 0.09 |  | 97
| 1 | 6 | 0.12333 | 0.13 | 0.1 | 0.1 | 0.1 | 0.1 | 0.13154 | 0.09 | 0.1 | 0.12 | 0.10206 | 0.1 | 0.11 | 0.1 | 0.11 | 0.09 | 0.128 | 0.09 | 0.11 | 0.12 | 0.09
| Mean |  | 0.13167 | 0.12667 | 0.1 | 0.096667 | 0.1 | 0.098333 | 0.13113 | 0.086667 | 0.1 | 0.12 | 0.10116 | 0.1 | 0.10833 | 0.096667 | 0.11333 | 0.09 | 0.12667 | 0.09 | 0.10833 | 0.11833 | 0.09
| Median |  | 0.13333 | 0.13 | 0.1 | 0.1 | 0.1 | 0.1 | 0.13136 | 0.0875 | 0.1 | 0.12 | 0.10094 | 0.1 | 0.11 | 0.1 | 0.11 | 0.09 | 0.1265 | 0.09 | 0.11 | 0.12 | 0.09
| Std Dev. |  | 0.0040825 | 0.005164 | 1.5202e-17 | 0.005164 | 1.5202e-17 | 0.0040825 | 0.001816 | 0.0040825 | 1.5202e-17 | 0 | 0.00065715 | 1.5202e-17 | 0.0040825 | 0.005164 | 0.005164 | 1.5202e-17 | 0.0031411 | 1.5202e-17 | 0.0040825 | 0.0075277 | 1.5202e-17
| Rel.Std.Dev. |  | 3.101% | 4.077% | 1.52e-14% | 5.342% | 1.52e-14% | 4.152% | 1.385% | 4.711% | 1.52e-14% | 0% | 0.6496% | 1.52e-14% | 3.768% | 5.342% | 4.556% | 1.689e-14% | 2.48% | 1.689e-14% | 3.768% | 6.361% | 1.689e-14%
| PDM3 |  | 23.77% | 19.07% | -5.996% | -9.13% | -5.996% | -7.563% | 23.26% | -18.53% | -5.996% | 12.8% | -4.902% | -5.996% | 1.837% | -9.13% | 6.537% | -15.4% | 19.07% | -15.4% | 1.837% | 11.24% | -15.4%
| Z-Score (Absolute) |  | 2.14 | 1.8 | 0 | 0.22 | 0 | 0.11 | 2.1 | 0.9 | 0 | 1.35 | 0.08 | 0 | 0.56 | 0.22 | 0.9 | 0.67 | 1.8 | 0.67 | 0.56 | 1.24 | 0.67
| Analytical results for Mn in OREAS 604 (Certified Value 231 ppm) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 0
| 1 | 1 | 237.814 | 234 | 225 | 222 | 227 | 221 | 243.9623 | 240 | 220 | 235 | 236.186 | 235 | 262.5 | 204 | 230 | 200 | 249.508 | 225 | 229 | 220 | 228 |  | 1
| 1 | 2 | 231.938 | 226 | 223 | 219 | 231 | 229 | 248.796 | 252 | 215 | 240 | 227.072 | 239 | 259.35 | 212 | 209 | 200 | 257.059 | 233 | 229 | 230 | 231 |  | 0
| 1 | 3 | 234.634 | 227 | 224 | 218 | 234 | 223 | 240.9184 | 254 | 425 | 242 | 226.027 | 236 | 261.45 | 213 | 221 | 200 | 265.511 | 232 | 234 | 230 | 238 |  | 16
| 1 | 4 | 232.338 | 231 | 225 | 219 | 229 | 223 | 248.4723 | 242 | 221 | 231 | 234.236 | 233 | 262.5 | 208 | 216 | 200 | 254.837 | 235 | 225 | 240 | 235 |  | 230.9176
| 1 | 5 | 228.544 | 232 | 227 | 220 | 234 | 226 | 237.9166 | 245 | 217 | 234 | 232.048 | 238 | 247.8 | 210 | 222 | 200 | 260.318 | 232 | 231 | 260 | 239 |  | 98
| 1 | 6 | 230.856 | 231 | 226 | 233 | 235 | 223 | 240.7888 | 249 | 219 | 237 | 230.053 | 238 | 256.2 | 207 | 225 | 200 | 264.071 | 238 | 235 | 240 | 233
| Mean |  | 232.6873 | 230.1667 | 225 | 221.8333 | 231.6667 | 224.1667 | 243.4757 | 247 | 252.8333 | 236.5 | 230.937 | 236.5 | 258.3 | 209 | 220.5 | 200 | 258.5507 | 232.5 | 230.5 | 236.6667 | 234
| Median |  | 232.138 | 231 | 225 | 219.5 | 232.5 | 223 | 242.4403 | 247 | 219.5 | 236 | 231.0505 | 237 | 260.4 | 209 | 221.5 | 200 | 258.6885 | 232.5 | 230 | 235 | 234
| Std Dev. |  | 3.2024 | 3.0605 | 1.4142 | 5.6362 | 3.2042 | 2.8577 | 4.4311 | 5.5857 | 84.3716 | 4.0373 | 3.9874 | 2.2583 | 5.6739 | 3.3466 | 7.287 | 0 | 5.9972 | 4.3243 | 3.6742 | 13.6626 | 4.1952
| Rel.Std.Dev. |  | 1.376% | 1.33% | 0.6285% | 2.541% | 1.383% | 1.275% | 1.82% | 2.261% | 33.37% | 1.707% | 1.727% | 0.9549% | 2.197% | 1.601% | 3.305% | 0% | 2.32% | 1.86% | 1.594% | 5.773% | 1.793%
| PDM3 |  | 0.7664% | -0.3252% | -2.563% | -3.934% | 0.3244% | -2.924% | 5.438% | 6.965% | 9.491% | 2.417% | 0.00838% | 2.417% | 11.86% | -9.492% | -4.511% | -13.39% | 11.97% | 0.6852% | -0.1809% | 2.49% | 1.335%
| Z-Score (Absolute) |  | 0.1 | 0.15 | 0.67 | 1.22 | 0 | 0.76 | 1.19 | 1.55 | 1.34 | 0.49 | 0.07 | 0.49 | 2.91 | 2.29 | 1.13 | 3.2 | 2.72 | 0.08 | 0.12 | 0.03 | 0.24
| Analytical results for Mo in OREAS 604 (Certified Value 4.02 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  | 3
| 1 | 1 | 4.07 | 4.2 | 4 | 6 | 4 | 4 | 4.2547 | 4 | 4 | 3 | 3.4183 | 4 | 4 | 2.3 | 6 | 4.268 | 3 | 4 | 4.6 | 4.8 |  |  | 1
| 1 | 2 | 3.975 | 4.1 | 4 | 6 | 4 | 4 | 4.0658 | 4 | 4 | 4 | 3.5221 | 4 | 4 | 1.87 | 6 | 4.048 | 3 | 4 | 4.8 | 4.8 |  |  | 0
| 1 | 3 | 4.1533 | 3.9 | 4 | 5 | 4 | 4 | 3.9367 | 4 | 4 | 4 | 3.4776 | 4 | 4 | 2.14 | 5 | 4.242 | 4 | 4 | 4.2 | 4.8 |  |  | 16
| 1 | 4 | 4.1167 | 4 | 4 | 6 | 4 | 5 | 4.1593 | 3 | 4 | 4 | 3.3749 | 4 | 4 | 3.06 | 6 | 3.977 | 4 | 4 | 4 | 4.8 |  |  | 4.0186
| 1 | 5 | 4.1 | 4.3 | 4 | 5 | 4 | 4 | 4.2539 | 4 | 4 | 3 | 4.8488 | 4 | 4 | 2.15 | 6 | 3.942 | 4 | 4 | 4.2 | 4.8 |  |  | 99
| 1 | 6 | 4.02 | 4.1 | 4 | 5 | 4 | 4 | 4.103 | 4 | 4 | 4 | 3.4189 | 4 | 4 | 2.03 | 5 | 4.156 | 4 | 4 | 4.2 | 4.8
| Mean |  | 4.0725 | 4.1 | 4 | 5.5 | 4 | 4.1667 | 4.1289 | 3.8333 | 4 | 3.6667 | 3.6768 | 4 | 4 | 2.2583 | 5.6667 | 4.1055 | 3.6667 | 4 | 4.3333 | 4.8
| Median |  | 4.085 | 4.1 | 4 | 5.5 | 4 | 4 | 4.1312 | 4 | 4 | 4 | 3.4482 | 4 | 4 | 2.145 | 6 | 4.102 | 4 | 4 | 4.2 | 4.8
| Std Dev. |  | 0.065589 | 0.14142 | 0 | 0.54772 | 0 | 0.40825 | 0.12162 | 0.40825 | 0 | 0.5164 | 0.5765 | 0 | 0 | 0.41787 | 0.5164 | 0.13721 | 0.5164 | 0 | 0.30111 | 0
| Rel.Std.Dev. |  | 1.611% | 3.449% | 0% | 9.959% | 0% | 9.798% | 2.946% | 10.65% | 0% | 14.08% | 15.68% | 0% | 0% | 18.5% | 9.113% | 3.342% | 14.08% | 0% | 6.949% | 0%
| PDM3 |  | 1.342% | 2.026% | -0.4625% | 36.86% | -0.4625% | 3.685% | 2.745% | -4.61% | -0.4625% | -8.757% | -8.506% | -0.4625% | -0.4625% | -43.8% | 41.01% | 2.163% | -8.757% | -0.4625% | 7.832% | 19.45%
| Z-Score (Absolute) |  | 0.33 | 0.46 | 0 | 6.84 | 0 | 0.76 | 0.59 | 0.76 | 0 | 1.52 | 2.54 | 0 | 0 | 8.68 | 7.6 | 0.48 | 1.52 | 0 | 1.52 | 3.65
| Analytical results for Na in OREAS 604 (Certified Value 461 ppm) | [23 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  |  | 0
| 1 | 1 | 366.6667 | 500 | 300 | 500 | 400 | 400 | 511.2802 | 500 | 600 | 433.762 | 410 | 500 | 400 | 500 | 400 | 560 | 600 | 500 | 500 |  |  |  | 1
| 1 | 2 | 366.6667 | 400 | 300 | 500 | 400 | 400 | 527.1035 | 400 | 600 | 438.353 | 420 | 500 | 400 | 500 | 400 | 590 | 600 | 500 | 500 |  |  |  | 0
| 1 | 3 | 366.6667 | 400 | 300 | 500 | 400 | 400 | 515.7269 | 400 | 600 | 436.498 | 410 | 500 | 400 | 500 | 400 | 590 | 600 | 500 | 600 |  |  |  | 16
| 1 | 4 | 333.3333 | 500 | 300 | 500 | 400 | 400 | 535.1771 | 400 | 600 | 441.399 | 430 | 500 | 400 | 500 | 400 | 560 | 600 | 400 | 600 |  |  |  | 461.4667
| 1 | 5 | 333.3333 | 400 | 300 | 500 | 400 | 400 | 514.0337 | 400 | 500 | 438.578 | > 100000 | 400 | 400 | 500 | 500 | 580 | 600 | 500 | 600 |  |  |  | 100
| 1 | 6 | 333.3333 | 400 | 300 | 500 | 400 | 400 | 523.0673 | 400 | 600 | 436.222 | 410 | 500 | 400 | 500 | 400 | 580 | 600 | 500 | 500
| Mean |  | 350 | 433.3333 | 300 | 500 | 400 | 400 | 521.0648 | 416.6667 | 583.3333 | 437.4687 | 416 | 483.3333 | 400 | 500 | 416.6667 | 576.6667 | 600 | 483.3333 | 550
| Median |  | 350 | 400 | 300 | 500 | 400 | 400 | 519.3971 | 400 | 600 | 437.4255 | 410 | 500 | 400 | 500 | 400 | 580 | 600 | 500 | 550
| Std Dev. |  | 18.2574 | 51.6398 | 0 | 0 | 0 | 0 | 9.0918 | 40.8248 | 40.8248 | 2.596 | 8.9443 | 40.8248 | 0 | 0 | 40.8248 | 13.6626 | 0 | 40.8248 | 54.7723
| Rel.Std.Dev. |  | 5.216% | 11.92% | 0% | 0% | 0% | 0% | 1.745% | 9.798% | 6.999% | 0.5934% | 2.15% | 8.447% | 0% | 0% | 9.798% | 2.369% | 0% | 8.447% | 9.959%
| PDM3 |  | -24.15% | -6.097% | -34.99% | 8.35% | -13.32% | -13.32% | 12.91% | -9.708% | 26.41% | -5.2% | -9.853% | 4.739% | -13.32% | 8.35% | -9.708% | 24.96% | 30.02% | 4.739% | 19.19%
| Z-Score (Absolute) |  | 1.29 | 0.06 | 2.02 | 0.92 | 0.55 | 0.55 | 1.23 | 0.31 | 2.14 | 0 | 0.32 | 0.67 | 0.55 | 0.92 | 0.31 | 2.05 | 2.39 | 0.67 | 1.65
| Analytical results for Nb in OREAS 604 (Certified Value 0.39 ppm) | [23 empty] | Certified
Nb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 14; 08: 15; 09: 17; 11: 19; 12: 25 |  |  |  | ppm
|  |  | 01: AR*OES; 02: AR*MS; 05: AR*OES/MS; 06: AR*OES; 07: AR*MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*OES/MS; 12: AR*MS |  |  |  | 2
|  |  | 01: 0.5g; 02: 40g; 05: 20g; 06: 0.15g; 07: 0.5g; 08: 01g; 09: 0.25g; 11: 0.25g; 12: 40g |  |  |  | 2
| 1 | 1 | 01: < 5; 02: 0.46; 05: 0.36839; 06: < 10; 07: 0.34; 08: 0.76; 09: 0.33; 11: 0.53; 12: 0.2 |  |  |  | 1
| 1 | 2 | 01: < 5; 02: 0.49; 05: 0.36677; 06: < 10; 07: 0.33; 08: 0.82; 09: 0.28; 11: 0.544; 12: < 0.2 |  |  |  | 0
| 1 | 3 | 01: < 5; 02: 0.43; 05: 0.35114; 06: < 10; 07: 0.33; 08: 0.77; 09: 0.31; 11: 0.56; 12: < 0.2 |  |  |  | 16
| 1 | 4 | 01: < 5; 02: 0.39; 05: 0.36468; 06: < 10; 07: 0.36; 08: 0.88; 09: 0.32; 11: 0.501; 12: < 0.2 |  |  |  | 0.38918
| 1 | 5 | 01: < 5; 02: 0.36; 05: 0.3484; 06: < 10; 07: 0.35; 08: 0.82; 09: 0.3; 11: 0.496; 12: 0.2 |  |  |  | 101
| 1 | 6 | 01: < 5; 02: 0.37; 05: 0.38733; 06: < 10; 07: 0.35; 08: 0.83; 09: 0.3; 11: 0.503; 12: < 0.2
| Mean |  | 02: 0.41667; 05: 0.36445; 07: 0.34333; 08: 0.81333; 09: 0.30667; 11: 0.52233; 12: 0.2
| Median |  | 02: 0.41; 05: 0.36572; 07: 0.345; 08: 0.82; 09: 0.305; 11: 0.5165; 12: 0.2
| Std Dev. |  | 02: 0.052026; 05: 0.013992; 07: 0.012111; 08: 0.043665; 09: 0.017512; 11: 0.026341; 12: 0
| Rel.Std.Dev. |  | 02: 12.49%; 05: 3.839%; 07: 3.527%; 08: 5.369%; 09: 5.71%; 11: 5.043%; 12: 0%
| PDM3 |  | 02: 7.062%; 05: -6.355%; 07: -11.78%; 08: 109%; 09: -21.2%; 11: 34.21%; 12: -48.61%
| Z-Score (Absolute) |  | 01: NA; 02: 0.61; 05: 0; 06: NA; 07: 0.25; 08: 5.24; 09: 0.67; 11: 1.75; 12: 1.92
| Analytical results for Nd in OREAS 604 (Certified Value 10.6 ppm) | [23 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  |  |  | 1
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  |  |  | 1
| 1 | 1 | 01: 11; 02: 12.2006; 05: 9.3311; 06: 10.29; 07: 12.13; 08: 8.46 |  |  |  | 1
| 1 | 2 | 01: 10.67; 02: 12.6289; 05: 9.0691; 06: 10.21; 07: 11.77; 08: 8.64 |  |  |  | 0
| 1 | 3 | 01: 10.94; 02: 12.3672; 05: 9.4855; 06: 10.06; 07: 12.66; 08: 8.6 |  |  |  | 16
| 1 | 4 | 01: 10.67; 02: 12.3228; 05: 9.1937; 06: 10.27; 07: 12.76; 08: 8.57 |  |  |  | 10.5923
| 1 | 5 | 01: 10.83; 02: 12.0785; 05: 9.5178; 06: 10.3; 07: 12.07; 08: 8.64 |  |  |  | 102
| 1 | 6 | 01: 10.96; 02: 12.4999; 05: 9.3595; 06: 10.33; 07: 11.75; 08: 8.69
| Mean |  | 01: 10.845; 02: 12.3497; 05: 9.3261; 06: 10.2433; 07: 12.19; 08: 8.6
| Median |  | 01: 10.885; 02: 12.345; 05: 9.3453; 06: 10.28; 07: 12.1; 08: 8.62
| Std Dev. |  | 01: 0.1468; 02: 0.19863; 05: 0.1714; 06: 0.098319; 07: 0.43216; 08: 0.07975
| Rel.Std.Dev. |  | 01: 1.354%; 02: 1.608%; 05: 1.838%; 06: 0.9598%; 07: 3.545%; 08: 0.9273%
| PDM3 |  | 01: 2.385%; 02: 16.59%; 05: -11.95%; 06: -3.295%; 07: 15.08%; 08: -18.81%
| Z-Score (Absolute) |  | 01: 0.14; 02: 0.85; 05: 0.57; 06: 0.14; 07: 0.78; 08: 0.92
| Analytical results for Ni in OREAS 604 (Certified Value 650 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 0
| 1 | 1 | 627.83 | 648.2 | 659 | 653 | 688 | 668 | 669.7623 | 663 | 676 | 672 | 573.569 | 640 | 657 | 605 | 550 | 600 | 644.313 | 621 | 681 | 593 | 642 |  | 1
| 1 | 2 | 626.685 | 655.2 | 648 | 641 | 687 | 665 | 680.8826 | 718 | 669 | 691 | 569.968 | 661 | 664 | 592 | 444 | 594 | 652.545 | 639 | 669 | 596 | 639 |  | 0
| 1 | 3 | 619.78 | 665.4 | 667 | 641 | 683 | 670 | 667.1784 | 697 | 660 | 692 | 566.477 | 635 | 659 | 624 | 487 | 574 | 657.699 | 636 | 671 | 620 | 647 |  | 16
| 1 | 4 | 613.52 | 644.1 | 657 | 648 | 679 | 666 | 685.1362 | 664 | 676 | 664 | 569.229 | 639 | 670 | 606 | 565 | 594 | 655.437 | 640 | 661 | 612 | 645 |  | 649.9844
| 1 | 5 | 614.67 | 653.3 | 672 | 640 | 678 | 680 | 681.4184 | 671 | 662 | 672 | 573.865 | 664 | 650 | 623 | 505 | 587 | 657.431 | 633 | 665 | 619 | 643 |  | 103
| 1 | 6 | 615.53 | 645.6 | 655 | 676 | 682 | 677 | 682.2324 | 672 | 678 | 683 | 566.922 | 635 | 649 | 610 | 448 | 585 | 664.375 | 649 | 691 | 616 | 641
| Mean |  | 619.6692 | 651.9667 | 659.6667 | 649.8333 | 682.8333 | 671 | 677.7684 | 680.8333 | 670.1667 | 679 | 570.005 | 645.6667 | 658.1667 | 610 | 499.8333 | 589 | 655.3 | 636.3333 | 673 | 609.3333 | 642.8333
| Median |  | 617.655 | 650.75 | 658 | 644.5 | 682.5 | 669 | 681.1505 | 671.5 | 672.5 | 677.5 | 569.5985 | 639.5 | 658 | 608 | 496 | 590.5 | 656.434 | 637.5 | 670 | 614 | 642.5
| Std Dev. |  | 6.2573 | 7.865 | 8.6178 | 13.7901 | 4.0702 | 6.1319 | 7.3957 | 21.9947 | 7.7567 | 11.4193 | 3.1672 | 13.2313 | 8.085 | 12.083 | 50.515 | 9.1214 | 6.6479 | 9.2448 | 11.0995 | 11.8603 | 2.8577
| Rel.Std.Dev. |  | 1.01% | 1.206% | 1.306% | 2.122% | 0.5961% | 0.9138% | 1.091% | 3.231% | 1.157% | 1.682% | 0.5557% | 2.049% | 1.228% | 1.981% | 10.11% | 1.549% | 1.014% | 1.453% | 1.649% | 1.946% | 0.4446%
| PDM3 |  | -4.664% | 0.305% | 1.49% | -0.02325% | 5.054% | 3.233% | 4.275% | 4.746% | 3.105% | 4.464% | -12.3% | -0.6643% | 1.259% | -6.152% | -23.1% | -9.382% | 0.8178% | -2.1% | 3.541% | -6.254% | -1.1%
| Z-Score (Absolute) |  | 1.04 | 0 | 0.25 | 0.24 | 0.99 | 0.61 | 0.83 | 0.69 | 0.58 | 0.87 | 2.63 | 0.2 | 0.2 | 1.35 | 4.88 | 2.02 | 0.11 | 0.5 | 0.67 | 1.37 | 0.29
| Analytical results for P in OREAS 604 (Certified Value 248 ppm) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 16 | 17 | 18 | 19 | 21 | 24 | 25 | [6 empty] | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | [6 empty] | 0
|  |  | 0.5g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 0.5g | 10g | 0.25g | 10g | 0.25g | 0.25g | 40g | 40g | [6 empty] | 0
| 1 | 1 | 250.34 | 230 | 220 | 230 | 220 | 301.3095 | 233 | 337 | 234 | 300 | 290 | < 100 | 350.171 | 363 | 260 | 240 | [6 empty] | 1
| 1 | 2 | 252.47 | 230 | 220 | 220 | 230 | 330.0197 | 220 | 361 | 234 | 200 | 240 | 200 | 406.987 | 407 | 280 | 240 | [6 empty] | 0
| 1 | 3 | 251.28 | 230 | 220 | 230 | 230 | 329.9459 | 224 | 361 | 229 | 300 | 280 | 100 | 337.604 | 307 | 260 | 240 | [6 empty] | 16
| 1 | 4 | 251.77 | 230 | 220 | 240 | 230 | 339.0126 | 225 | 347 | 231 | 300 | 270 | 200 | 382.09 | 463 | 280 | 240 | [6 empty] | 247.7249
| 1 | 5 | 248.43 | 230 | 220 | 230 | 240 | 285.6176 | 223 | 334 | 233 | 300 | 250 | 200 | 380.337 | 444 | 280 | 240 | [6 empty] | 104
| 1 | 6 | 248.51 | 230 | 230 | 240 | 220 | 278.6906 | 222 | 365 | 231 | 200 | 240 | 100 | 340.295 | 427 | 300 | 240
| Mean |  | 250.4667 | 230 | 221.6667 | 231.6667 | 228.3333 | 310.766 | 224.5 | 350.8333 | 232 | 266.6667 | 261.6667 | 160 | 366.2473 | 401.8333 | 276.6667 | 240
| Median |  | 250.81 | 230 | 220 | 230 | 230 | 315.6277 | 223.5 | 354 | 232 | 300 | 260 | 200 | 365.254 | 417 | 280 | 240
| Std Dev. |  | 1.6951 | 0 | 4.0825 | 7.5277 | 7.5277 | 25.6404 | 4.5056 | 13.3928 | 2 | 51.6398 | 21.3698 | 54.7723 | 27.7908 | 57.7596 | 15.0555 | 0
| Rel.Std.Dev. |  | 0.6768% | 0% | 1.842% | 3.249% | 3.297% | 8.251% | 2.007% | 3.817% | 0.8621% | 19.36% | 8.167% | 34.23% | 7.588% | 14.37% | 5.442% | 0%
| PDM3 |  | 1.107% | -7.155% | -10.52% | -6.482% | -7.828% | 25.45% | -9.375% | 41.62% | -6.348% | 7.646% | 5.628% | -35.41% | 47.84% | 62.21% | 11.68% | -3.118%
| Z-Score (Absolute) |  | 0.16 | 0.47 | 0.72 | 0.42 | 0.52 | 2.01 | 0.69 | 3.25 | 0.41 | 0.66 | 0.51 | 2.62 | 3.72 | 4.81 | 0.97 | 0.16
| Analytical results for Pb in OREAS 604 (Certified Value 703 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 0
| 1 | 1 | 680.02 | 700 | 702 | 682 | 720 | 680 | 699.1033 | 730 | 742 | 743 | 714.2872 | 712.9 | 734 | 667 | 783 | 629 | 743.42 | 626 | 718 | 680 | 668 |  | 1
| 1 | 2 | 686.44 | 701.7 | 699 | 664 | 721 | 700 | 708.8561 | 750 | 745 | 769 | 720.9529 | 721.6 | 738 | 675 | 623 | 629 | 754.21 | 647 | 706 | 684 | 662 |  | 0
| 1 | 3 | 686.09 | 717.1 | 699 | 666 | 722 | 698 | 705.2415 | 760 | 743 | 760 | 716.1656 | 719.5 | 737 | 697 | 761 | 605 | 739.19 | 638 | 715 | 662 | 662 |  | 16
| 1 | 4 | 673.44 | 719.1 | 702 | 672 | 718 | 698 | 712.3502 | 730 | 748 | 744 | 709.3592 | 699.6 | 756 | 683 | 718 | 638 | 769.08 | 638 | 695 | 670 | 655 |  | 703.095
| 1 | 5 | 680 | 711.5 | 705 | 669 | 716 | 713 | 698.1327 | 760 | 738 | 748 | 719.1205 | 728.8 | 728 | 686 | 742 | 629 | 765.52 | 641 | 709 | 683 | 658 |  | 105
| 1 | 6 | 669 | 699 | 698 | 708 | 719 | 700 | 717.6215 | 750 | 738 | 755 | 706.0457 | 707.6 | 742 | 683 | 693 | 620 | 737.32 | 648 | 722 | 690 | 657
| Mean |  | 679.165 | 708.0667 | 700.8333 | 676.8333 | 719.3333 | 698.1667 | 706.8842 | 746.6667 | 742.3333 | 753.1667 | 714.3219 | 715 | 739.1667 | 681.8333 | 720 | 625 | 751.4567 | 639.6667 | 710.8333 | 678.1667 | 660.3333
| Median |  | 680.01 | 706.6 | 700.5 | 670.5 | 719.5 | 699 | 707.0488 | 750 | 742.5 | 751.5 | 715.2264 | 716.2 | 737.5 | 683 | 730 | 629 | 748.815 | 639.5 | 712 | 681.5 | 660
| Std Dev. |  | 6.9052 | 8.977 | 2.6394 | 16.5217 | 2.1602 | 10.553 | 7.6011 | 13.6626 | 3.9328 | 10.1472 | 5.7187 | 10.4936 | 9.4745 | 10.1669 | 57.0544 | 11.3314 | 13.6444 | 7.9666 | 9.704 | 10.2843 | 4.6762
| Rel.Std.Dev. |  | 1.017% | 1.268% | 0.3766% | 2.441% | 0.3003% | 1.512% | 1.075% | 1.83% | 0.5298% | 1.347% | 0.8006% | 1.468% | 1.282% | 1.491% | 7.924% | 1.813% | 1.816% | 1.245% | 1.365% | 1.516% | 0.7082%
| PDM3 |  | -3.404% | 0.7071% | -0.3217% | -3.735% | 2.31% | -0.7009% | 0.5389% | 6.197% | 5.581% | 7.122% | 1.597% | 1.693% | 5.13% | -3.024% | 2.404% | -11.11% | 6.878% | -9.021% | 1.101% | -3.546% | -6.082%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.17 | 0.87 | 0.26 | 0.23 | 0.03 | 0.9 | 0.8 | 1.05 | 0.15 | 0.16 | 0.73 | 0.61 | 0.28 | 1.84 | 1.01 | 1.6 | 0.06 | 0.7 | 1.11
| Analytical results for Pd in OREAS 604 (Indicative Value 5.01 ppm) | [23 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab |  | 1
| No | No | 01: 09; 02: 14; 05: 25 |  | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 05: AR*MS |  | 2
|  |  | 01: 20g; 02: 0.5g; 05: 40g |  | 2
| 1 | 1 | 01: 4.7668; 02: 5.197; 05: 4.89 |  | 1
| 1 | 2 | 01: 4.8595; 02: 5.378; 05: 5.03 |  | 1
| 1 | 3 | 01: 4.7978; 02: 5.248; 05: 4.92 |  | 16
| 1 | 4 | 01: 4.9216; 02: 5.139; 05: 4.95 |  | 5.0124
| 1 | 5 | 01: 4.7803; 02: 5.295; 05: 4.94 |  | 14
| 1 | 6 | 01: 4.8683; 02: 5.221; 05: 5.02
| Mean |  | 01: 4.8324; 02: 5.2463; 05: 4.9583
| Median |  | 01: 4.8287; 02: 5.2345; 05: 4.945
| Std Dev. |  | 01: 0.060308; 02: 0.082829; 05: 0.055648
| Rel.Std.Dev. |  | 01: 1.248%; 02: 1.579%; 05: 1.122%
| PDM3 |  | 01: -3.59%; 02: 4.668%; 05: -1.078%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.54; 05: 0
| Analytical results for Pr in OREAS 604 (Certified Value 2.83 ppm) | [23 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  | 2
| 1 | 1 | 01: 3.14; 02: 3.1309; 05: 2.7173; 06: 2.65; 07: 3.41; 08: 2.05 |  | 1
| 1 | 2 | 01: 3.02; 02: 3.2694; 05: 2.5687; 06: 2.64; 07: 3.24; 08: 2.12 |  | 0
| 1 | 3 | 01: 3; 02: 3.1772; 05: 2.6087; 06: 2.65; 07: 3.46; 08: 2.07 |  | 16
| 1 | 4 | 01: 3.04; 02: 3.1297; 05: 2.4776; 06: 2.66; 07: 3.5; 08: 2.09 |  | 2.8257
| 1 | 5 | 01: 3.03; 02: 3.0711; 05: 2.6749; 06: 2.72; 07: 3.26; 08: 2.12 |  | 106
| 1 | 6 | 01: 3.08; 02: 3.1962; 05: 2.6429; 06: 2.68; 07: 3.3; 08: 2.13
| Mean |  | 01: 3.0517; 02: 3.1624; 05: 2.615; 06: 2.6667; 07: 3.3617; 08: 2.0967
| Median |  | 01: 3.035; 02: 3.154; 05: 2.6258; 06: 2.655; 07: 3.355; 08: 2.105
| Std Dev. |  | 01: 0.050761; 02: 0.068115; 05: 0.084736; 06: 0.029439; 07: 0.10962; 08: 0.032042
| Rel.Std.Dev. |  | 01: 1.663%; 02: 2.154%; 05: 3.24%; 06: 1.104%; 07: 3.261%; 08: 1.528%
| PDM3 |  | 01: 7.998%; 02: 11.92%; 05: -7.456%; 06: -5.628%; 07: 18.97%; 08: -25.8%
| Z-Score (Absolute) |  | 01: 0.47; 02: 0.75; 05: 0.6; 06: 0.47; 07: 1.24; 08: 1.88
| Analytical results for Pt in OREAS 604 (Indicative Value 890 ppb) | [23 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab; 05: Lab |  | 1
| No | No | 01: 09; 02: 14; 05: 25 |  | ppb
|  |  | 01: AR*OES/MS; 02: AR*MS; 05: AR*MS |  | 0
|  |  | 01: 20g; 02: 0.5g; 05: 40g |  | 0
| 1 | 1 | 01: 834.4027; 02: 885; 05: 925 |  | 1
| 1 | 2 | 01: 859.2292; 02: 908; 05: 920 |  | 1
| 1 | 3 | 01: 850.5562; 02: 905; 05: 925 |  | 16
| 1 | 4 | 01: 854.9694; 02: 874; 05: 925 |  | 890.2983
| 1 | 5 | 01: 849.206; 02: 920; 05: 915 |  | 12
| 1 | 6 | 01: 861.0058; 02: 899; 05: 915
| Mean |  | 01: 851.5615; 02: 898.5; 05: 920.8333
| Median |  | 01: 852.7628; 02: 902; 05: 922.5
| Std Dev. |  | 01: 9.5978; 02: 16.5982; 05: 4.916
| Rel.Std.Dev. |  | 01: 1.127%; 02: 1.847%; 05: 0.5339%
| PDM3 |  | 01: -4.351%; 02: 0.9212%; 05: 3.43%
| Z-Score (Absolute) |  | 01: 1.42; 02: 0; 05: 0.67
| Analytical results for Rb in OREAS 604 (Certified Value 8.39 ppm) | [23 empty] | Certified
Rb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 10: 24; 11: 25 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 10: 40g; 11: 40g |  | 2
| 1 | 1 | 01: 10.81; 02: 11.624; 05: 8.2825; 06: 7.34; 07: 11.4; 08: 11.8; 09: 7.29; 10: 7.8; 11: 6.8 |  | 1
| 1 | 2 | 01: 10.68; 02: 11.5436; 05: 8.0139; 06: 7.43; 07: 12.62; 08: 9.4; 09: 7.58; 10: 7.4; 11: 7.05 |  | 0
| 1 | 3 | 01: 10.05; 02: 11.5126; 05: 8.6491; 06: 7.37; 07: 12.13; 08: 10.2; 09: 7.83; 10: 7.2; 11: 7 |  | 16
| 1 | 4 | 01: 10.96; 02: 11.8391; 05: 8.5023; 06: 7.48; 07: 12.41; 08: 9.8; 09: 7.51; 10: 7.3; 11: 7.15 |  | 8.3917
| 1 | 5 | 01: 10.78; 02: 11.3983; 05: 8.2094; 06: 7.31; 07: 11.28; 08: 10.6; 09: 7.56; 10: 7.4; 11: 7.2 |  | 107
| 1 | 6 | 01: 10.53; 02: 11.9437; 05: 8.144; 06: 7.61; 07: 11.42; 08: 9.2; 09: 7.48; 10: 7.6; 11: 7.45
| Mean |  | 01: 10.635; 02: 11.6436; 05: 8.3002; 06: 7.4233; 07: 11.8767; 08: 10.1667; 09: 7.5417; 10: 7.45; 11: 7.1083
| Median |  | 01: 10.73; 02: 11.5838; 05: 8.246; 06: 7.4; 07: 11.775; 08: 10; 09: 7.535; 10: 7.4; 11: 7.1
| Std Dev. |  | 01: 0.32017; 02: 0.20785; 05: 0.23558; 06: 0.11021; 07: 0.58188; 08: 0.95009; 09: 0.17498; 10: 0.21679; 11: 0.21775
| Rel.Std.Dev. |  | 01: 3.011%; 02: 1.785%; 05: 2.838%; 06: 1.485%; 07: 4.899%; 08: 9.345%; 09: 2.32%; 10: 2.91%; 11: 3.063%
| PDM3 |  | 01: 26.73%; 02: 38.75%; 05: -1.091%; 06: -11.54%; 07: 41.53%; 08: 21.15%; 09: -10.13%; 10: -11.22%; 11: -15.29%
| Z-Score (Absolute) |  | 01: 1.39; 02: 1.89; 05: 0; 06: 0.5; 07: 2.02; 08: 1.06; 09: 0.43; 10: 0.48; 11: 0.67
| Analytical results for Re in OREAS 604 (Indicative Value 1 ppb) | [23 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 14; 07: 15 |  | ppb
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*OES; 06: AR*MS; 07: AR*OES/MS |  | 0
|  |  | 01: 40g; 02: 20g; 05: 0.15g; 06: 0.5g; 07: 01g |  | 0
| 1 | 1 | 01: 2; 02: < 1; 05: < 5000; 06: < 1; 07: 10 |  | 1
| 1 | 2 | 01: 2; 02: < 1; 05: < 5000; 06: 2; 07: 10 |  | 4
| 1 | 3 | 01: 2; 02: < 1; 05: < 5000; 06: < 1; 07: < 10 |  | 16
| 1 | 4 | 01: 1; 02: < 1; 05: < 5000; 06: 1; 07: 10 |  | 1.25
| 1 | 5 | 01: 1; 02: < 1; 05: < 5000; 06: 1; 07: 20 |  | 13
| 1 | 6 | 01: 1; 02: < 1; 05: < 5000; 06: 1; 07: 10
| Mean |  | 01: 1.5; 06: 1.25; 07: 12
| Median |  | 01: 1.5; 06: 1; 07: 10
| Std Dev. |  | 01: 0.54772; 06: 0.5; 07: 4.4721
| Rel.Std.Dev. |  | 01: 36.51%; 06: 40%; 07: 37.27%
| PDM3 |  | 01: 20%; 06: 0%; 07: 860%
| Z-Score (Absolute) |  | 01: 0; 02: NA; 05: NA; 06: 0.67; 07: 10.11
| Analytical results for Ru in OREAS 604 (Indicative Value < 2 ppb) | [23 empty] | Indicative
Ru | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 14 |  | ppb
|  |  | 01: AR*MS |  | 0
|  |  | 01: 0.5g |  | 0
| 1 | 1 | 01: < 2 |  | 1
| 1 | 2 | 01: < 2 |  | 8
| 1 | 3 | 01: < 2 |  | 16
| 1 | 4 | 01: < 2 |  | < 2
| 1 | 5 | 01: < 2 |  | 14
| 1 | 6 | 01: < 2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for S in OREAS 604 (Certified Value 4.26 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 19 | 23 | 24 | 25 |  |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  |  |  | 3
| 1 | 1 | 4.5333 | 3.7074 | 4.38 | 4.19 | 4.23 | 4.23 | 4.4766 | 4.53 | 4.34 | 4.4659 | 4.0999 | 4.149 | 4 | - | 3.978 | 4.06 | 4.47 | 4.17 |  |  |  |  | 1
| 1 | 2 | 4.5333 | 3.805 | 4.34 | 4.08 | 4.24 | 4.3 | 4.5374 | 4.65 | 4.27 | 4.4077 | 4.2521 | 4.148 | 3.91 | - | 4.083 | 4.03 | 4.98 | 4.08 |  |  |  |  | 0
| 1 | 3 | 4.53 | 3.9574 | 4.36 | 4.08 | 4.39 | 4.33 | 4.5298 | 4.7 | 4.23 | 4.3674 | 4.0744 | 4.168 | 4.07 | - | 4 | 4.08 | 4.28 | 4.16 |  |  |  |  | 16
| 1 | 4 | 4.4767 | 4.0882 | 4.4 | 4.12 | 4.35 | 4.3 | 4.5348 | 4.48 | 4.28 | 4.3844 | 4.0981 | 4.28 | 3.99 | - | 4.049 | 3.94 | 4.63 | 4.17 |  |  |  |  | 4.258
| 1 | 5 | 4.4633 | 3.93 | 4.39 | 4.11 | 4.35 | 4.38 | 4.5517 | 4.57 | 4.27 | 4.4077 | 4.2955 | 4.088 | 4.08 | - | 4.086 | 4.06 | 4.86 | 4.2 |  |  |  |  | 108
| 1 | 6 | 4.38 | 3.9638 | 4.38 | 4.36 | 4.38 | 4.35 | 4.5651 | 4.56 | 4.28 | 4.3779 | 4.1203 | 4.09 | 3.98 | - | 4.054 | 4.13 | 4.3 | 4.13
| Mean |  | 4.4861 | 3.9086 | 4.375 | 4.1567 | 4.3233 | 4.315 | 4.5326 | 4.5817 | 4.2783 | 4.4018 | 4.1567 | 4.1538 | 4.005 |  | 4.0417 | 4.05 | 4.5867 | 4.1517
| Median |  | 4.5033 | 3.9437 | 4.38 | 4.115 | 4.35 | 4.315 | 4.5361 | 4.565 | 4.275 | 4.396 | 4.1101 | 4.1485 | 3.995 |  | 4.0515 | 4.06 | 4.55 | 4.165
| Std Dev. |  | 0.060421 | 0.13368 | 0.021679 | 0.10746 | 0.070333 | 0.051672 | 0.030298 | 0.080353 | 0.035449 | 0.035329 | 0.092871 | 0.070121 | 0.062849 |  | 0.043976 | 0.063246 | 0.29022 | 0.041673
| Rel.Std.Dev. |  | 1.347% | 3.42% | 0.4955% | 2.585% | 1.627% | 1.197% | 0.6685% | 1.754% | 0.8286% | 0.8026% | 2.234% | 1.688% | 1.569% |  | 1.088% | 1.562% | 6.327% | 1.004%
| PDM3 |  | 5.358% | -8.204% | 2.749% | -2.379% | 1.535% | 1.34% | 6.449% | 7.602% | 0.4784% | 3.379% | -2.378% | -2.446% | -5.941% |  | -5.08% | -4.884% | 7.72% | -2.496%
| Z-Score (Absolute) |  | 0.86 | 1.54 | 0.4 | 0.67 | 0.19 | 0.15 | 1.06 | 1.26 | 0 | 0.51 | 0.51 | 0.52 | 1.14 | NA | 0.98 | 0.95 | 0.95 | 0.53
| Analytical results for Sb in OREAS 604 (Certified Value 126 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  | 0
| 1 | 1 | 87.2153 | 114.47 | 141 | 106 | 126 | 158 | 91.3267 | 120 | 127 | 125 | 104.978 | 141.8 | 159.17 | 121 | 59.5 | 135 | 100.481 | 67 | 149 | 138 |  |  | 1
| 1 | 2 | 86.9773 | 119.32 | 139 | 103 | 130 | 160 | 88.7795 | 130 | 125 | 132 | 104.7451 | 139.8 | 159.67 | 120 | 53.5 | 142 | 100.55 | 71 | 150 | 143 |  |  | 0
| 1 | 3 | 87.8313 | 117.85 | 139 | 103 | 147 | 163 | 88.7123 | 130 | 124 | 125 | 99.9243 | 137.96 | 159.58 | 123 | 60 | 136 | 103.494 | 73 | 150 | 137 |  |  | 16
| 1 | 4 | 86.9307 | 122.47 | 142 | 104 | 147 | 160 | 90.0538 | 130 | 124 | 123 | 106.8358 | 144.38 | 162.74 | 126 | 55 | 143 | 100.847 | 73 | 147 | 135 |  |  | 126.2324
| 1 | 5 | 86.1163 | 118.82 | 143 | 105 | 148 | 161 | 85.7927 | 130 | 123 | 119 | 101.415 | 143.59 | 157.04 | 128 | 57.5 | 138 | 103.234 | 68 | 154 | 134 |  |  | 109
| 1 | 6 | 85.5447 | 114.46 | 141 | 102 | 152 | 162 | 88.2374 | 120 | 125 | 122 | 103.563 | 145.99 | 148.33 | 123 | 57 | 133 | 99.764 | 68 | 152 | 138
| Mean |  | 86.7693 | 117.8983 | 140.8333 | 103.8333 | 141.6667 | 160.6667 | 88.817 | 126.6667 | 124.6667 | 124.3333 | 103.5769 | 142.2533 | 157.755 | 123.5 | 57.0833 | 137.8333 | 101.395 | 70 | 150.3333 | 137.5
| Median |  | 86.954 | 118.335 | 141 | 103.5 | 147 | 160.5 | 88.7459 | 130 | 124.5 | 124 | 104.154 | 142.695 | 159.375 | 123 | 57.25 | 137 | 100.6985 | 69.5 | 150 | 137.5
| Std Dev. |  | 0.81514 | 3.0773 | 1.6021 | 1.472 | 10.8197 | 1.7512 | 1.8613 | 5.164 | 1.3663 | 4.3665 | 2.5284 | 2.9976 | 4.9637 | 3.0166 | 2.5183 | 3.9707 | 1.5683 | 2.6833 | 2.4221 | 3.1464
| Rel.Std.Dev. |  | 0.9394% | 2.61% | 1.138% | 1.418% | 7.637% | 1.09% | 2.096% | 4.077% | 1.096% | 3.512% | 2.441% | 2.107% | 3.146% | 2.443% | 4.412% | 2.881% | 1.547% | 3.833% | 1.611% | 2.288%
| PDM3 |  | -31.26% | -6.602% | 11.57% | -17.74% | 12.23% | 27.28% | -29.64% | 0.344% | -1.24% | -1.504% | -17.95% | 12.69% | 24.97% | -2.165% | -54.78% | 9.19% | -19.68% | -44.55% | 19.09% | 8.926%
| Z-Score (Absolute) |  | 1.32 | 0.23 | 0.57 | 0.73 | 0.6 | 1.27 | 1.25 | 0.08 | 0.01 | 0.01 | 0.73 | 0.62 | 1.23 | 0.04 | 2.37 | 0.47 | 0.81 | 1.91 | 0.91 | 0.46
| Analytical results for Sc in OREAS 604 (Certified Value < 2 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.5g | 40g | 40g |  |  |  | 2
| 1 | 1 | 1.3367 | 1.4 | 1 | 1 | 1 | 1 | 1.6833 | 1 | 1 | 1.1406 | 1.15 | 1.58 | 0.8 | 1.8 | < 3 | 1.37 | < 5 | 2 | 1 |  |  |  | 1
| 1 | 2 | 1.34 | 1.4 | 1 | 1 | 1 | 1 | 1.6977 | 2 | 1 | 1.1583 | 1.24 | 1.67 | 0.9 | 1.5 | < 3 | 1.44 | < 5 | 2 | 1 |  |  |  | 0
| 1 | 3 | 1.33 | 1.3 | 1 | 1 | 1 | 1 | 1.6216 | 1 | 1 | 1.1022 | 1.2 | 1.75 | 0.9 | 2 | < 3 | 1.43 | < 5 | 2 | 1 |  |  |  | 16
| 1 | 4 | 1.33 | 1.4 | 1 | 1 | 1 | 1 | 1.767 | 2 | 1 | 1.126 | 1.2 | 1.6 | 0.9 | 1.9 | < 3 | 1.41 | < 5 | 2 | 1 |  |  |  | < 2
| 1 | 5 | 1.3233 | 1.4 | 1 | 1 | 1 | 1 | 1.6839 | 2 | 1 | 1.1065 | 1.19 | 1.56 | 0.9 | 1.6 | < 3 | 1.45 | < 5 | 2 | 1 |  |  |  | 110
| 1 | 6 | 1.2933 | 1.4 | 1 | 1 | 1 | 1 | 1.7037 | 2 | 1 | 1.133 | 1.2 | 1.53 | 0.8 | 1.8 | < 3 | 1.4 | < 5 | 2 | 1
| Mean |  | 1.3256 | 1.3833 | 1 | 1 | 1 | 1 | 1.6928 | 1.6667 | 1 | 1.1278 | 1.1967 | 1.615 | 0.86667 | 1.7667 |  | 1.4167 |  | 2 | 1
| Median |  | 1.33 | 1.4 | 1 | 1 | 1 | 1 | 1.6908 | 2 | 1 | 1.1295 | 1.2 | 1.59 | 0.9 | 1.8 |  | 1.42 |  | 2 | 1
| Std Dev. |  | 0.016822 | 0.040825 | 0 | 0 | 0 | 0 | 0.046665 | 0.5164 | 0 | 0.021145 | 0.028752 | 0.081179 | 0.05164 | 0.18619 |  | 0.029439 |  | 0 | 0
| Rel.Std.Dev. |  | 1.269% | 2.951% | 0% | 0% | 0% | 0% | 2.757% | 30.98% | 0% | 1.875% | 2.403% | 5.027% | 5.958% | 10.54% |  | 2.078% |  | 0% | 0%
PDM3
| Z-Score (Absolute) |  | 0.44 | 0.64 | 0.67 | 0.67 | 0.67 | 0.67 | 1.7 | 1.61 | 0.67 | 0.24 | 0 | 1.43 | 1.13 | 1.95 | NA | 0.75 | NA | 2.75 | 0.67
| Analytical results for Se in OREAS 604 (Certified Value 45 ppm) | [23 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 14 | 15 | 17 | 18 | 19 | 21 | 24 | 25 | [10 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | [10 empty] | 1
|  |  | 0.5g | 40g | 20g | 0.15g | 0.5g | 01g | 0.25g | 10g | 0.25g | 0.25g | 40g | 40g | [10 empty] | 1
| 1 | 1 | 50.0633 | 45 | 46.1741 | 45 | 42 | 44.89 | 55 | 35 | 36.062 | 47 | 52.5 | 40 | [10 empty] | 1
| 1 | 2 | 50.39 | 43 | 46.4097 | 43 | 42.7 | 46.52 | 46 | 33 | 36.532 | 49 | 52.5 | 41 | [10 empty] | 0
| 1 | 3 | 50.8067 | 43 | 46.1705 | 44 | 41.8 | 46.43 | 53 | 44 | 36.885 | 51 | 48 | 41 | [10 empty] | 16
| 1 | 4 | 49.6133 | 44 | 45.0592 | 44 | 41.4 | 46.29 | 58 | 42 | 34.934 | 51 | 48 | 41 | [10 empty] | 45.0449
| 1 | 5 | 49.73 | 43 | 45.8711 | 43 | 43.3 | 44.85 | 52 | 46 | 36.578 | 51 | 49 | 40 | [10 empty] | 111
| 1 | 6 | 49.895 | 45 | 47.1622 | 43 | 42.5 | 45.86 | 49 | 39 | 35.856 | 50 | 49.5 | 42
| Mean |  | 50.0831 | 43.8333 | 46.1411 | 43.6667 | 42.2833 | 45.8067 | 52.1667 | 39.8333 | 36.1412 | 49.8333 | 49.9167 | 40.8333
| Median |  | 49.9792 | 43.5 | 46.1723 | 43.5 | 42.25 | 46.075 | 52.5 | 40.5 | 36.297 | 50.5 | 49.25 | 41
| Std Dev. |  | 0.44696 | 0.98319 | 0.68702 | 0.8165 | 0.68532 | 0.76018 | 4.2622 | 5.1153 | 0.69881 | 1.6021 | 2.0837 | 0.75277
| Rel.Std.Dev. |  | 0.8924% | 2.243% | 1.489% | 1.87% | 1.621% | 1.66% | 8.17% | 12.84% | 1.934% | 3.215% | 4.174% | 1.844%
| PDM3 |  | 11.18% | -2.69% | 2.434% | -3.06% | -6.131% | 1.691% | 15.81% | -11.57% | -19.77% | 10.63% | 10.82% | -9.35%
| Z-Score (Absolute) |  | 0.79 | 0.15 | 0.2 | 0.17 | 0.38 | 0.15 | 1.1 | 0.75 | 1.3 | 0.75 | 0.77 | 0.6
| Analytical results for Sm in OREAS 604 (Certified Value 1.94 ppm) | [23 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 40g |  |  |  | 2
| 1 | 1 | 01: 2.11; 02: 2.2062; 05: 1.8302; 06: 1.81; 07: 2.16; 08: 1.53 |  |  |  | 1
| 1 | 2 | 01: 2.03; 02: 2.1876; 05: 1.8117; 06: 1.78; 07: 2.16; 08: 1.56 |  |  |  | 0
| 1 | 3 | 01: 2.02; 02: 2.1709; 05: 1.8331; 06: 1.75; 07: 2.28; 08: 1.53 |  |  |  | 16
| 1 | 4 | 01: 2; 02: 2.2161; 05: 1.8316; 06: 1.77; 07: 2.44; 08: 1.52 |  |  |  | 1.9443
| 1 | 5 | 01: 2.08; 02: 2.1356; 05: 1.8279; 06: 1.86; 07: 2.22; 08: 1.6 |  |  |  | 112
| 1 | 6 | 01: 2.09; 02: 2.3169; 05: 1.7885; 06: 1.84; 07: 2.08; 08: 1.62
| Mean |  | 01: 2.055; 02: 2.2056; 05: 1.8205; 06: 1.8017; 07: 2.2233; 08: 1.56
| Median |  | 01: 2.055; 02: 2.1969; 05: 1.8291; 06: 1.795; 07: 2.19; 08: 1.545
| Std Dev. |  | 01: 0.044159; 02: 0.061536; 05: 0.017524; 06: 0.042622; 07: 0.12549; 08: 0.041473
| Rel.Std.Dev. |  | 01: 2.149%; 02: 2.79%; 05: 0.9626%; 06: 2.366%; 07: 5.644%; 08: 2.659%
| PDM3 |  | 01: 5.691%; 02: 13.43%; 05: -6.37%; 06: -7.338%; 07: 14.35%; 08: -19.77%
| Z-Score (Absolute) |  | 01: 0.39; 02: 0.89; 05: 0.39; 06: 0.45; 07: 0.95; 08: 1.26
| Analytical results for Sn in OREAS 604 (Certified Value 2.81 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 14 | 15 | 16 | 17 | 18 | 19 | 24 | 25 | [10 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 40g | 40g | [10 empty] | 3
| 1 | 1 | < 5 | 2.91 | 2.9466 | < 10 | 3 | 3.15 | < 10 | 3.1 | < 10 | 2.444 | 3 | 2.7 | [10 empty] | 1
| 1 | 2 | < 5 | 2.69 | 2.94 | < 10 | 3 | 3.07 | < 10 | 2.5 | < 10 | 2.409 | 3 | 2.9 | [10 empty] | 0
| 1 | 3 | < 5 | 2.63 | 2.8907 | < 10 | 3 | 3.16 | < 10 | 2.6 | < 10 | 2.406 | 3 | 2.8 | [10 empty] | 16
| 1 | 4 | < 5 | 2.66 | 3.0027 | < 10 | 3 | 3.28 | < 10 | 2.5 | < 10 | 2.317 | 3 | 2.7 | [10 empty] | 2.8083
| 1 | 5 | < 5 | 2.58 | 2.8756 | < 10 | 3 | 2.79 | < 10 | 2.7 | < 10 | 2.374 | 3 | 2.7 | [10 empty] | 113
| 1 | 6 | < 5 | 2.77 | 2.9206 | < 10 | 3 | 3.11 | < 10 | 2.4 | < 10 | 2.368 | 3 | 2.7
| Mean |  | 02: 2.7067; 05: 2.9294; 07: 3; 08: 3.0933; 11: 2.6333; 13: 2.3863; 14: 3; 15: 2.75
| Median |  | 02: 2.675; 05: 2.9303; 07: 3; 08: 3.13; 11: 2.55; 13: 2.39; 14: 3; 15: 2.7
| Std Dev. |  | 02: 0.1181; 05: 0.045293; 07: 0; 08: 0.16452; 11: 0.25033; 13: 0.043638; 14: 0; 15: 0.083666
| Rel.Std.Dev. |  | 02: 4.363%; 05: 1.546%; 07: 0%; 08: 5.319%; 11: 9.506%; 13: 1.829%; 14: 0%; 15: 3.042%
| PDM3 |  | 02: -3.619%; 05: 4.311%; 07: 6.826%; 08: 10.15%; 11: -6.23%; 13: -15.03%; 14: 6.826%; 15: -2.076%
| Z-Score (Absolute) |  | NA | 0.56 | 0.38 | NA | 0.67 | 1.32 | NA | 1.26 | NA | 1.91 | 0.67 | 0.38
| Analytical results for Sr in OREAS 604 (Certified Value 37.7 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  |  | 1
| 1 | 1 | 30.471 | 50.75 | 41 | 39 | 39 | 40 | 52.5279 | 50 | 40 | 33.2697 | 34.47 | 28.66 | 33.8 | 41.9 | 33 | 37.93 | 37 | 29 | 22 | 16 |  |  | 1
| 1 | 2 | 29.855 | 49.46 | 40 | 37 | 40 | 42 | 53.8537 | 49 | 41 | 32.6356 | 35.09 | 29.77 | 35.1 | 39.9 | 33 | 39.51 | 37 | 30 | 22 | 17.3 |  |  | 0
| 1 | 3 | 29.8807 | 49.9 | 41 | 38 | 38 | 42 | 54.1282 | 49 | 42 | 30.9149 | 34.98 | 29.95 | 37.6 | 41.1 | 33 | 41.82 | 37 | 29 | 21 | 16.2 |  |  | 16
| 1 | 4 | 29.5867 | 48.46 | 40 | 37 | 39 | 41 | 54.5144 | 50 | 41 | 32.205 | 34.45 | 26.91 | 37.7 | 40.8 | 34 | 37.95 | 38 | 28 | 21 | 16.4 |  |  | 37.7147
| 1 | 5 | 29.5983 | 48.42 | 40 | 40 | 38 | 43 | 52.5281 | 49 | 40 | 33.4135 | 35.13 | 28.74 | 34.8 | 41.9 | 34 | 40.27 | 37 | 30 | 22 | 16.7 |  |  | 114
| 1 | 6 | 29.4506 | 50.14 | 40 | 37 | 39 | 41 | 54.2896 | 51 | 40 | 32.7429 | 35.44 | 28.21 | 37.6 | 41.4 | 33 | 39.87 | 37 | 30 | 23 | 17.8
| Mean |  | 29.8071 | 49.5217 | 40.3333 | 38 | 38.8333 | 41.5 | 53.6403 | 49.6667 | 40.6667 | 32.5303 | 34.9267 | 28.7067 | 36.1 | 41.1667 | 33.3333 | 39.5583 | 37.1667 | 29.3333 | 21.8333 | 16.7333
| Median |  | 29.7267 | 49.68 | 40 | 37.5 | 39 | 41.5 | 53.9909 | 49.5 | 40.5 | 32.6893 | 35.035 | 28.7 | 36.35 | 41.25 | 33 | 39.69 | 37 | 29.5 | 22 | 16.55
| Std Dev. |  | 0.36543 | 0.93583 | 0.5164 | 1.2649 | 0.75277 | 1.0488 | 0.88807 | 0.8165 | 0.8165 | 0.90536 | 0.39246 | 1.1095 | 1.7344 | 0.75807 | 0.5164 | 1.4805 | 0.40825 | 0.8165 | 0.75277 | 0.69186
| Rel.Std.Dev. |  | 1.226% | 1.89% | 1.28% | 3.329% | 1.938% | 2.527% | 1.656% | 1.644% | 2.008% | 2.783% | 1.124% | 3.865% | 4.804% | 1.841% | 1.549% | 3.743% | 1.098% | 2.784% | 3.448% | 4.135%
| PDM3 |  | -20.97% | 31.31% | 6.943% | 0.7564% | 2.966% | 10.04% | 42.23% | 31.69% | 7.827% | -13.75% | -7.393% | -23.88% | -4.281% | 9.153% | -11.62% | 4.888% | -1.453% | -22.22% | -42.11% | -55.63%
| Z-Score (Absolute) |  | 1.28 | 1.97 | 0.45 | 0.07 | 0.21 | 0.65 | 2.65 | 2 | 0.51 | 0.83 | 0.44 | 1.47 | 0.24 | 0.59 | 0.7 | 0.33 | 0.07 | 1.36 | 2.6 | 3.44
| Analytical results for Ta in OREAS 604 (Indicative Value < 0.01 ppm) | [23 empty] | Indicative
Ta | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 14; 06: 15; 07: 17; 08: 19; 09: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*MS |  |  | 3
|  |  | 01: 40g; 02: 20g; 05: 0.5g; 06: 01g; 07: 0.25g; 08: 0.25g; 09: 40g |  |  | 3
| 1 | 1 | 01: 0.09; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.1; 08: < 0.05; 09: < 0.005 |  |  | 1
| 1 | 2 | 01: 0.09; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.1; 08: < 0.05; 09: < 0.005 |  |  | 6
| 1 | 3 | 01: 0.09; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.12; 08: < 0.05; 09: < 0.005 |  |  | 16
| 1 | 4 | 01: 0.1; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.11; 08: < 0.05; 09: < 0.005 |  |  | < 0.01
| 1 | 5 | 01: 0.08; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.1; 08: < 0.05; 09: < 0.005 |  |  | 12
| 1 | 6 | 01: 0.09; 02: < 0.01; 05: < 0.01; 06: < 0.1; 07: 0.09; 08: < 0.05; 09: < 0.005
| Mean |  | 01: 0.09; 07: 0.10333
| Median |  | 01: 0.09; 07: 0.1
| Std Dev. |  | 01: 0.0063246; 07: 0.010328
| Rel.Std.Dev. |  | 01: 7.027%; 07: 9.995%
PDM3
| Z-Score (Absolute) |  | 01: 0.67; 02: NA; 05: NA; 06: NA; 07: 0.67; 08: NA; 09: NA
| Analytical results for Tb in OREAS 604 (Certified Value 0.19 ppm) | [23 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 11: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*MS |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 11: 40g |  |  | 2
| 1 | 1 | 01: 0.2; 02: 0.22793; 05: 0.18779; 06: 0.194; 07: 0.23; 08: 0.16; 09: 0.179; 11: 0.135 |  |  | 1
| 1 | 2 | 01: 0.2; 02: 0.24147; 05: 0.17719; 06: 0.191; 07: 0.21; 08: 0.13; 09: 0.176; 11: 0.14 |  |  | 0
| 1 | 3 | 01: 0.19; 02: 0.23462; 05: 0.171; 06: 0.191; 07: 0.22; 08: 0.15; 09: 0.175; 11: 0.14 |  |  | 16
| 1 | 4 | 01: 0.2; 02: 0.24067; 05: 0.15917; 06: 0.186; 07: 0.24; 08: 0.16; 09: 0.17; 11: 0.14 |  |  | 0.18568
| 1 | 5 | 01: 0.2; 02: 0.23061; 05: 0.18155; 06: 0.19; 07: 0.24; 08: 0.15; 09: 0.176; 11: 0.145 |  |  | 115
| 1 | 6 | 01: 0.2; 02: 0.23722; 05: 0.17548; 06: 0.192; 07: 0.2; 08: 0.13; 09: 0.168; 11: 0.15
| Mean |  | 01: 0.19833; 02: 0.23542; 05: 0.17536; 06: 0.19067; 07: 0.22333; 08: 0.14667; 09: 0.174; 11: 0.14167
| Median |  | 01: 0.2; 02: 0.23592; 05: 0.17634; 06: 0.191; 07: 0.225; 08: 0.15; 09: 0.1755; 11: 0.14
| Std Dev. |  | 01: 0.0040825; 02: 0.0054291; 05: 0.0097708; 06: 0.0026583; 07: 0.01633; 08: 0.013663; 09: 0.0041473; 11: 0.005164
| Rel.Std.Dev. |  | 01: 2.058%; 02: 2.306%; 05: 5.572%; 06: 1.394%; 07: 7.312%; 08: 9.315%; 09: 2.383%; 11: 3.645%
| PDM3 |  | 01: 6.814%; 02: 26.79%; 05: -5.556%; 06: 2.685%; 07: 20.28%; 08: -21.01%; 09: -6.291%; 11: -23.7%
| Z-Score (Absolute) |  | 01: 0.4; 02: 1.37; 05: 0.2; 06: 0.2; 07: 1.05; 08: 0.95; 09: 0.24; 11: 1.08
| Analytical results for Te in OREAS 604 (Certified Value 25.3 ppm) | [23 empty] | Certified
Te | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 14; 07: 15; 08: 17; 09: 19; 11: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*OES; 06: AR*MS; 07: AR*OES; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*MS |  |  | 1
|  |  | 01: 40g; 02: 20g; 05: 0.15g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 11: 40g |  |  | 2
| 1 | 1 | 01: 29.92; 02: 23.0099; 05: 22; 06: 25.17; 07: 27; 08: 22.9; 09: 25.591; 11: 25.6 |  |  | 1
| 1 | 2 | 01: 27.57; 02: 23.5376; 05: 25; 06: 26.44; 07: 24; 08: 17.6; 09: 25.368; 11: 26 |  |  | 0
| 1 | 3 | 01: 26.77; 02: 22.9401; 05: 24; 06: 25.36; 07: 26; 08: 17.5; 09: 25.31; 11: 26 |  |  | 16
| 1 | 4 | 01: 27.82; 02: 23.2291; 05: 25; 06: 24.98; 07: 24; 08: 20.4; 09: 25.194; 11: 25.7 |  |  | 25.3294
| 1 | 5 | 01: 28; 02: 22.5632; 05: 20; 06: 25.99; 07: 27; 08: 19.3; 09: 25.046; 11: 25.6 |  |  | 116
| 1 | 6 | 01: 26.41; 02: 23.1383; 05: 27; 06: 25.85; 07: 28; 08: 16.3; 09: 24.828; 11: 25.9
| Mean |  | 01: 27.7483; 02: 23.0697; 05: 23.8333; 06: 25.6317; 07: 26; 08: 19; 09: 25.2228; 11: 25.8
| Median |  | 01: 27.695; 02: 23.0741; 05: 24.5; 06: 25.605; 07: 26.5; 08: 18.45; 09: 25.252; 11: 25.8
| Std Dev. |  | 01: 1.23; 02: 0.32435; 05: 2.4833; 06: 0.55517; 07: 1.6733; 08: 2.3983; 09: 0.26541; 11: 0.18974
| Rel.Std.Dev. |  | 01: 4.433%; 02: 1.406%; 05: 10.42%; 06: 2.166%; 07: 6.436%; 08: 12.62%; 09: 1.052%; 11: 0.7354%
| PDM3 |  | 01: 9.55%; 02: -8.921%; 05: -5.906%; 06: 1.193%; 07: 2.647%; 08: -24.99%; 09: -0.4208%; 11: 1.858%
| Z-Score (Absolute) |  | 01: 1.44; 02: 1.47; 05: 0.99; 06: 0.13; 07: 0.36; 08: 4; 09: 0.13; 11: 0.23
| Analytical results for Th in OREAS 604 (Certified Value 3.74 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 08 | 09 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | [7 empty] | 2
|  |  | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 0.5g | 40g | 40g | [7 empty] | 2
| 1 | 1 | 4.19 | < 20 | < 20 | < 20 | < 20 | 3.6364 | 3.1026 | 3.19 | 4.4 | 4 | < 10 | 3.718 | 5 | 1.9 | 2.34 | [7 empty] | 1
| 1 | 2 | 4.1 | < 20 | < 20 | < 20 | < 20 | 3.7487 | 3.4665 | 3.18 | 4.58 | 3.2 | < 10 | 3.784 | 5 | 1.8 | 2.42 | [7 empty] | 0
| 1 | 3 | 4.13 | < 20 | < 20 | < 20 | < 20 | 3.7565 | 3.5172 | 3.16 | 4.49 | 3.4 | < 10 | 3.953 | 5 | 1.65 | 2.36 | [7 empty] | 16
| 1 | 4 | 4.21 | < 20 | < 20 | < 20 | < 20 | 3.8479 | 3.3762 | 3.16 | 4.56 | 3.7 | < 10 | 3.755 | 4 | 1.7 | 2.46 | [7 empty] | 3.7373
| 1 | 5 | 4.25 | < 20 | < 20 | < 20 | < 20 | 3.6518 | 3.5489 | 3.25 | 4.32 | 3.7 | < 10 | 3.773 | 5 | 1.75 | 2.42 | [7 empty] | 117
| 1 | 6 | 4.39 | < 20 | < 20 | < 20 | < 20 | 3.8146 | 3.623 | 3.2 | 4.2 | 3.1 | < 10 | 3.798 | 5 | 1.8 | 2.48
| Mean |  | 4.2117 |  |  |  |  | 3.7427 | 3.4391 | 3.19 | 4.425 | 3.5167 |  | 3.7968 | 4.8333 | 1.7667 | 2.4133
| Median |  | 4.2 |  |  |  |  | 3.7526 | 3.4919 | 3.185 | 4.445 | 3.55 |  | 3.7785 | 5 | 1.775 | 2.42
| Std Dev. |  | 0.10284 |  |  |  |  | 0.084851 | 0.18436 | 0.033466 | 0.14748 | 0.34303 |  | 0.081347 | 0.40825 | 0.08756 | 0.05465
| Rel.Std.Dev. |  | 2.442% |  |  |  |  | 2.267% | 5.361% | 1.049% | 3.333% | 9.754% |  | 2.143% | 8.447% | 4.956% | 2.265%
| PDM3 |  | 12.69% |  |  |  |  | 0.1429% | -7.98% | -14.64% | 18.4% | -5.904% |  | 1.592% | 29.33% | -52.73% | -35.43%
| Z-Score (Absolute) |  | 0.77 | NA | NA | NA | NA | 0.15 | 0.16 | 0.58 | 1.05 | 0.15 | NA | 0.22 | 1.59 | 2.46 | 1.61
| Analytical results for Ti in OREAS 604 (Certified Value 97 ppm) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  |  | 0
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  |  |  | 0
| 1 | 1 | 105 | 105 | 100 | 100 | 100 | 100 | 134.4109 | 100 | < 100 | 66 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60 | < 50 |  |  |  |  | 1
| 1 | 2 | 105 | 101 | 100 | 100 | 100 | 100 | 136.3409 | < 100 | < 100 | 67 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60 | < 50 |  |  |  |  | 0
| 1 | 3 | 105 | 101 | 100 | 100 | 100 | 100 | 134.5814 | < 100 | < 100 | 65 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60 | < 50 |  |  |  |  | 16
| 1 | 4 | 105 | 100 | 100 | 100 | 100 | 100 | 139.1672 | < 100 | < 100 | 65 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60 | < 50 |  |  |  |  | 97.0729
| 1 | 5 | 105 | 97 | 100 | 100 | 100 | 100 | 130.3748 | < 100 | < 100 | 68 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60 | < 50 |  |  |  |  | 118
| 1 | 6 | 105 | 101 | 100 | 100 | 100 | 100 | 137.9341 | < 100 | < 100 | 68 | 100 | < 100 | < 100 | < 100 | < 100 | < 100 | 60 | < 50
| Mean |  | 105 | 100.8333 | 100 | 100 | 100 | 100 | 135.4682 | 100 |  | 66.5 | 100 | [5 empty] | 60
| Median |  | 105 | 101 | 100 | 100 | 100 | 100 | 135.4611 | 100 |  | 66.5 | 100 | [5 empty] | 60
| Std Dev. |  | 0 | 2.5626 | 0 | 0 | 0 | 0 | 3.1103 |  |  | 1.3784 | 0 | [5 empty] | 0
| Rel.Std.Dev. |  | 0% | 2.541% | 0% | 0% | 0% | 0% | 2.296% |  |  | 2.073% | 0% | [5 empty] | 0%
| PDM3 |  | 8.166% | 3.874% | 3.015% | 3.015% | 3.015% | 3.015% | 39.55% | 3.015% |  | -31.49% | 3.015% | [5 empty] | -38.19%
| Z-Score (Absolute) | [16 empty] | NA | NA | NA | bar: 02 to 25
| Analytical results for Tl in OREAS 604 (Certified Value 7.27 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*MS | [7 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g | [7 empty] | 2
| 1 | 1 | < 5 | 8.52 | 6.7889 | 8 | < 10 | 5.4149 | 7.63 | 9.02 | 8.72 | < 10 | 7.589 | 8 | 7 | 8.4 | 6.8 | [7 empty] | 1
| 1 | 2 | < 5 | 8.4 | 6.8254 | 6 | < 10 | 5.3191 | 7.75 | 9.42 | 7.2 | < 10 | 7.704 | 4 | < 5 | 9 | 6.9 | [7 empty] | 0
| 1 | 3 | < 5 | 8.49 | 6.6604 | 7 | < 10 | 5.2014 | 7.65 | 9.38 | 7.13 | < 10 | 7.819 | < 2 | < 5 | 7.7 | 6.9 | [7 empty] | 16
| 1 | 4 | < 5 | 8.39 | 6.9531 | 8 | < 10 | 5.1648 | 7.53 | 9.29 | 7.83 | < 10 | 7.518 | < 2 | < 5 | 7.9 | 6.8 | [7 empty] | 7.2658
| 1 | 5 | < 5 | 8.52 | 6.7378 | 6 | < 10 | 5.2181 | 7.88 | 8.99 | 8.16 | < 10 | 7.637 | < 2 | 6 | 8.3 | 6.8 | [7 empty] | 119
| 1 | 6 | < 5 | 8.52 | 6.9515 | 8 | < 10 | 5.4675 | 7.56 | 8.49 | 6.7 | < 10 | 7.6 | < 2 | 6 | 8.1 | 6.8
| Mean |  |  | 8.4733 | 6.8195 | 7.1667 |  | 5.2976 | 7.6667 | 9.0983 | 7.6233 |  | 7.6445 | 6 | 6.3333 | 8.2333 | 6.8333
| Median |  |  | 8.505 | 6.8071 | 7.5 |  | 5.2686 | 7.64 | 9.155 | 7.515 |  | 7.6185 | 6 | 6 | 8.2 | 6.8
| Std Dev. |  |  | 0.06186 | 0.11679 | 0.98319 |  | 0.12355 | 0.12972 | 0.34856 | 0.74947 |  | 0.10499 | 2.8284 | 0.57735 | 0.45461 | 0.05164
| Rel.Std.Dev. |  |  | 0.7301% | 1.713% | 13.72% |  | 2.332% | 1.692% | 3.831% | 9.831% |  | 1.373% | 47.14% | 9.116% | 5.522% | 0.7557%
| PDM3 |  |  | 16.62% | -6.143% | -1.365% |  | -27.09% | 5.517% | 25.22% | 4.92% |  | 5.212% | -17.42% | -12.83% | 13.32% | -5.952%
| Z-Score (Absolute) |  | NA | 1.03 | 0.55 | 0.22 | NA | 2 | 0.26 | 1.62 | 0.22 | NA | 0.24 | 1.33 | 1.01 | 0.8 | 0.54
| Analytical results for Tm in OREAS 604 (Indicative Value 31 ppb) | [23 empty] | Indicative
Tm | Round | Replicate | 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | [7 empty] | 1
| No | No | 02: 02; 05: 09; 06: 14; 07: 15; 08: 25 | [7 empty] | ppb
|  |  | 02: AR*MS; 05: AR*OES/MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS | [7 empty] | 0
|  |  | 02: 40g; 05: 20g; 06: 0.5g; 07: 01g; 08: 40g | [7 empty] | 0
| 1 | 1 | 02: 40; 05: 36.3165; 06: 30; 07: < 100; 08: 25 | [7 empty] | 1
| 1 | 2 | 02: 30; 05: 36.4; 06: 30; 07: < 100; 08: 25 | [7 empty] | 7
| 1 | 3 | 02: 30; 05: 34.1137; 06: 30; 07: < 100; 08: 25 | [7 empty] | 16
| 1 | 4 | 02: 30; 05: 32.5712; 06: 30; 07: < 100; 08: 20 | [7 empty] | 31.1221
| 1 | 5 | 02: 40; 05: 37.4861; 06: 30; 07: < 100; 08: 25 | [7 empty] | 13
| 1 | 6 | 02: 30; 05: 36.4586; 06: 30; 07: < 100; 08: 25
| Mean |  | 02: 33.3333; 05: 35.5577; 06: 30; 08: 24.1667
| Median |  | 02: 30; 05: 36.3583; 06: 30; 08: 25
| Std Dev. |  | 02: 5.164; 05: 1.8341; 06: 0; 08: 2.0412
| Rel.Std.Dev. |  | 02: 15.49%; 05: 5.158%; 06: 0%; 08: 8.447%
| PDM3 |  | 02: 7.105%; 05: 14.25%; 06: -3.605%; 08: -22.35%
| Z-Score (Absolute) |  | 02: 0.37; 05: 0.98; 06: 0.37; 07: NA; 08: 1.46
| Analytical results for U in OREAS 604 (Certified Value 1.22 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 05 | 07 | 08 | 09 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 24 | 25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [8 empty] | 2
|  |  | 40g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 40g | 40g | [8 empty] | 2
| 1 | 1 | 1.41 | < 10 | < 10 | < 10 | 1.2638 | < 10 | 1.0861 | 1.16 | 1.46 | 1.5 | < 10 | 1.252 | 1.05 | 0.92 | [8 empty] | 1
| 1 | 2 | 1.38 | < 10 | < 10 | < 10 | 1.252 | < 10 | 1.0596 | 1.13 | 1.44 | 1.19 | < 10 | 1.295 | 1.05 | 0.92 | [8 empty] | 0
| 1 | 3 | 1.36 | < 10 | < 10 | < 10 | 1.2779 | < 10 | 1.0178 | 1.15 | 1.47 | 1.29 | < 10 | 1.354 | 1.05 | 0.92 | [8 empty] | 16
| 1 | 4 | 1.41 | < 10 | < 10 | < 10 | 1.2892 | < 10 | 1.042 | 1.11 | 1.46 | 1.23 | < 10 | 1.339 | 1.05 | 0.94 | [8 empty] | 1.215
| 1 | 5 | 1.45 | < 10 | < 10 | < 10 | 1.2643 | < 10 | 1.0928 | 1.15 | 1.4 | 1.38 | < 10 | 1.291 | 1.1 | 0.96 | [8 empty] | 120
| 1 | 6 | 1.45 | < 10 | < 10 | < 10 | 1.3157 | < 10 | 1.0553 | 1.13 | 1.31 | 1.16 | < 10 | 1.308 | 1.1 | 0.98
| Mean |  | 1.41 |  |  |  | 1.2771 |  | 1.0589 | 1.1383 | 1.4233 | 1.2917 |  | 1.3065 | 1.0667 | 0.94
| Median |  | 1.41 |  |  |  | 1.2711 |  | 1.0575 | 1.14 | 1.45 | 1.26 |  | 1.3015 | 1.05 | 0.93
| Std Dev. |  | 0.036332 |  |  |  | 0.022822 |  | 0.027838 | 0.018348 | 0.060882 | 0.12859 |  | 0.03649 | 0.02582 | 0.025298
| Rel.Std.Dev. |  | 2.577% |  |  |  | 1.787% |  | 2.629% | 1.612% | 4.277% | 9.956% |  | 2.793% | 2.421% | 2.691%
| PDM3 |  | 16.05% |  |  |  | 5.112% |  | -12.85% | -6.312% | 17.14% | 6.308% |  | 7.528% | -12.21% | -22.64%
| Z-Score (Absolute) |  | 0.65 | NA | NA | NA | 0 | NA | 1.06 | 0.67 | 0.82 | 0.07 | NA | 0.14 | 1.02 | 1.64
| Analytical results for V in OREAS 604 (Certified Value 10.2 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 1
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g |  |  |  | 1
| 1 | 1 | 12.4933 | 11 | 10 | 9 | 9 | 9 | 12.6236 | 9 | 10 | 8.9462 | 10 | 10.77 | 11 | 14 | 8 | 10.606 | 9 | 8 | 10 |  |  |  | 1
| 1 | 2 | 12.1033 | 11 | 10 | 9 | 9 | 9 | 12.8463 | 9 | 10 | 8.7073 | 10 | 10.5 | 11 | 13 | 8 | 10.744 | 8 | 8 | 8 |  |  |  | 0
| 1 | 3 | 12.1767 | 11 | 10 | 9 | 10 | 9 | 12.6953 | 10 | 10 | 8.3373 | 10 | 11.98 | 12 | 13 | 8 | 11.201 | 9 | 9 | 10 |  |  |  | 16
| 1 | 4 | 12.09 | 12 | 10 | 9 | 10 | 9 | 13.0258 | 9 | 10 | 8.4647 | 10 | 10.96 | 11 | 14 | 8 | 11.092 | 9 | 8 | 10 |  |  |  | 10.1531
| 1 | 5 | 12.0833 | 12 | 10 | 9 | 9 | 10 | 12.5813 | 9 | 10 | 8.7012 | 10 | 10.38 | 11 | 14 | 8 | 11.378 | 9 | 9 | 10 |  |  |  | 121
| 1 | 6 | 11.44 | 11 | 10 | 9 | 9 | 9 | 12.9196 | 9 | 10 | 9.1032 | 10 | 10.12 | 11 | 14 | 8 | 10.637 | 9 | 9 | 8
| Mean |  | 12.0644 | 11.3333 | 10 | 9 | 9.3333 | 9.1667 | 12.782 | 9.1667 | 10 | 8.71 | 10 | 10.785 | 11.1667 | 13.6667 | 8 | 10.943 | 8.8333 | 8.5 | 9.3333
| Median |  | 12.0967 | 11 | 10 | 9 | 9 | 9 | 12.7708 | 9 | 10 | 8.7043 | 10 | 10.635 | 11 | 14 | 8 | 10.918 | 9 | 8.5 | 10
| Std Dev. |  | 0.34322 | 0.5164 | 0 | 0 | 0.5164 | 0.40825 | 0.17627 | 0.40825 | 0 | 0.28621 | 0 | 0.65519 | 0.40825 | 0.5164 | 0 | 0.32397 | 0.40825 | 0.54772 | 1.0328
| Rel.Std.Dev. |  | 2.845% | 4.556% | 0% | 0% | 5.533% | 4.454% | 1.379% | 4.454% | 0% | 3.286% | 0% | 6.075% | 3.656% | 3.779% | 0% | 2.961% | 4.622% | 6.444% | 11.07%
| PDM3 |  | 18.82% | 11.62% | -1.508% | -11.36% | -8.074% | -9.716% | 25.89% | -9.716% | -1.508% | -14.21% | -1.508% | 6.223% | 9.983% | 34.61% | -21.21% | 7.78% | -13% | -16.28% | -8.074%
| Z-Score (Absolute) |  | 1.48 | 0.9 | 0 | 0.67 | 0.45 | 0.56 | 1.88 | 0.56 | 0 | 0.87 | 0 | 0.53 | 0.79 | 2.47 | 1.35 | 0.64 | 0.79 | 1.01 | 0.45
| Analytical results for W in OREAS 604 (Certified Value 3.5 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 08 | 09 | 11 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 25 | [7 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 40g | 01g | 01g | 01g | 20g | 0.15g | 01g | 0.5g | 01g | 0.25g | 10g | 0.25g | 0.5g | 40g | [7 empty] | 2
| 1 | 1 | < 5 | 4.43 | < 10 | < 10 | < 10 | 3.1556 | < 10 | 3.2169 | 3.55 | 4.74 | 3.8 | < 10 | 4.37 | < 2 | 2.6 | [7 empty] | 1
| 1 | 2 | < 5 | 4.31 | < 10 | < 10 | < 10 | 3.1039 | < 10 | 3.2503 | 3.42 | 4.7 | 3.3 | < 10 | 4.375 | 3 | 2.6 | [7 empty] | 0
| 1 | 3 | < 5 | 3.72 | < 10 | < 10 | < 10 | 3.1158 | < 10 | 3.1325 | 3.45 | 4.43 | 3.2 | < 10 | 4.37 | < 2 | 2.6 | [7 empty] | 16
| 1 | 4 | < 5 | 4.17 | < 10 | < 10 | < 10 | 3.1564 | < 10 | 3.2084 | 3.56 | 4.65 | 3.8 | < 10 | 4.17 | 2 | 2.5 | [7 empty] | 3.497
| 1 | 5 | < 5 | 3.66 | < 10 | < 10 | < 10 | 3.0582 | < 10 | 3.1456 | 3.63 | 5.02 | 3.8 | < 10 | 4.187 | 3 | 2.4 | [7 empty] | 122
| 1 | 6 | < 5 | 3.68 | < 10 | < 10 | < 10 | 3.1849 | < 10 | 2.9998 | 3.57 | 4.52 | 3.1 | < 10 | 4.328 | < 2 | 2.4
| Mean |  | 02: 3.995; 08: 3.1291; 11: 3.1589; 12: 3.53; 13: 4.6767; 14: 3.5; 16: 4.3; 17: 2.6667; 18: 2.5167
| Median |  | 02: 3.945; 08: 3.1357; 11: 3.177; 12: 3.555; 13: 4.675; 14: 3.55; 16: 4.349; 17: 3; 18: 2.55
| Std Dev. |  | 02: 0.34818; 08: 0.045616; 11: 0.089846; 12: 0.079246; 13: 0.20422; 14: 0.33466; 16: 0.095789; 17: 0.57735; 18: 0.098319
| Rel.Std.Dev. |  | 02: 8.715%; 08: 1.458%; 11: 2.844%; 12: 2.245%; 13: 4.367%; 14: 9.562%; 16: 2.228%; 17: 21.65%; 18: 3.907%
| PDM3 |  | 02: 14.24%; 08: -10.52%; 11: -9.668%; 12: 0.9436%; 13: 33.73%; 14: 0.08568%; 16: 22.96%; 17: -23.74%; 18: -28.03%
| Z-Score (Absolute) |  | NA | 0.67 | NA | NA | NA | 0.51 | NA | 0.46 | 0.04 | 1.6 | 0 | NA | 1.09 | 1.14 | 1.34
| Analytical results for Y in OREAS 604 (Certified Value 3.45 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 25 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [9 empty] | 2
|  |  | 0.5g | 40g | 20g | 0.15g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 40g | [9 empty] | 2
| 1 | 1 | 4.3306 | 3.54 | 3.7419 | 3 | 3.8623 | 3.57 | 4 | < 5 | 2.72 | 1 | 3.454 | 3 | 2.97 | [9 empty] | 1
| 1 | 2 | 4.3072 | 3.42 | 3.815 | 3 | 3.6817 | 3.64 | 4 | < 5 | 2.23 | 2 | 3.484 | 3 | 3.05 | [9 empty] | 0
| 1 | 3 | 4.3206 | 3.33 | 3.7124 | 3 | 3.8365 | 3.66 | 4 | < 5 | 2.61 | 2 | 3.525 | 3 | 3.04 | [9 empty] | 16
| 1 | 4 | 4.3731 | 3.45 | 3.8041 | 3 | 3.61 | 3.55 | 4 | < 5 | 2.49 | 2 | 3.405 | 3 | 3.04 | [9 empty] | 3.447
| 1 | 5 | 4.1963 | 3.33 | 3.7161 | 3 | 3.8611 | 3.55 | 4 | < 5 | 2.57 | 2 | 3.506 | 3 | 3.08 | [9 empty] | 123
| 1 | 6 | 4.3502 | 3.48 | 3.8735 | 3 | 3.8061 | 3.6 | 4 | < 5 | 2.48 | 1 | 3.412 | 3 | 3.12
| Mean |  | 4.313 | 3.425 | 3.7772 | 3 | 3.7763 | 3.595 | 4 |  | 2.5167 | 1.6667 | 3.4643 | 3 | 3.05
| Median |  | 4.3256 | 3.435 | 3.773 | 3 | 3.8213 | 3.585 | 4 |  | 2.53 | 2 | 3.469 | 3 | 3.045
| Std Dev. |  | 0.06167 | 0.083606 | 0.06418 | 0 | 0.10555 | 0.046797 | 0 |  | 0.16561 | 0.5164 | 0.049334 | 0 | 0.0498
| Rel.Std.Dev. |  | 1.43% | 2.441% | 1.699% | 0% | 2.795% | 1.302% | 0% |  | 6.581% | 30.98% | 1.424% | 0% | 1.633%
| PDM3 |  | 25.12% | -0.6394% | 9.577% | -12.97% | 9.551% | 4.292% | 16.04% |  | -26.99% | -51.65% | 0.5017% | -12.97% | -11.52%
| Z-Score (Absolute) |  | 1.4 | 0.03 | 0.53 | 0.71 | 0.53 | 0.24 | 0.89 | NA | 1.49 | 2.86 | 0.03 | 0.71 | 0.63
| Analytical results for Yb in OREAS 604 (Certified Value 0.2 ppm) | [23 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 14; 07: 15; 08: 17; 09: 19; 11: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS; 11: AR*MS |  |  |  | 2
|  |  | 01: 40g; 02: 20g; 05: 01g; 06: 0.5g; 07: 01g; 08: 0.25g; 09: 0.25g; 11: 40g |  |  |  | 3
| 1 | 1 | 01: 0.21; 02: 0.20666; 05: 0.18153; 06: 0.21; 07: 0.19; 08: 0.3; 09: 0.194; 11: 0.18 |  |  |  | 1
| 1 | 2 | 01: 0.21; 02: 0.21123; 05: 0.19917; 06: 0.2; 07: 0.2; 08: 0.2; 09: 0.202; 11: 0.18 |  |  |  | 0
| 1 | 3 | 01: 0.2; 02: 0.20337; 05: 0.19212; 06: 0.2; 07: 0.23; 08: 0.2; 09: 0.198; 11: 0.16 |  |  |  | 16
| 1 | 4 | 01: 0.23; 02: 0.22724; 05: 0.16899; 06: 0.2; 07: 0.14; 08: 0.2; 09: 0.196; 11: 0.17 |  |  |  | 0.19529
| 1 | 5 | 01: 0.21; 02: 0.21925; 05: 0.20232; 06: 0.2; 07: 0.15; 08: 0.2; 09: 0.194; 11: 0.17 |  |  |  | 124
| 1 | 6 | 01: 0.2; 02: 0.19958; 05: 0.19251; 06: 0.2; 07: 0.18; 08: 0.2; 09: 0.196; 11: 0.17
| Mean |  | 01: 0.21; 02: 0.21122; 05: 0.18944; 06: 0.20167; 07: 0.18167; 08: 0.21667; 09: 0.19667; 11: 0.17167
| Median |  | 01: 0.21; 02: 0.20894; 05: 0.19231; 06: 0.2; 07: 0.185; 08: 0.2; 09: 0.196; 11: 0.17
| Std Dev. |  | 01: 0.010954; 02: 0.010381; 05: 0.012313; 06: 0.0040825; 07: 0.033116; 08: 0.040825; 09: 0.0030111; 11: 0.0075277
| Rel.Std.Dev. |  | 01: 5.216%; 02: 4.915%; 05: 6.499%; 06: 2.024%; 07: 18.23%; 08: 18.84%; 09: 1.531%; 11: 4.385%
| PDM3 |  | 01: 7.532%; 02: 8.157%; 05: -2.996%; 06: 3.265%; 07: -6.976%; 08: 10.95%; 09: 0.7045%; 11: -12.1%
| Z-Score (Absolute) |  | 01: 0.77; 02: 0.85; 05: 0.58; 06: 0.22; 07: 1.09; 08: 0.11; 09: 0.11; 11: 1.75
| Analytical results for Zn in OREAS 604 (Certified Value 0.254 wt.%) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 40g | 01g | 01g | 01g | 01g | 20g | 01g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 0.25g | 0.5g | 40g | 40g |  | 3
| 1 | 1 | 0.26564 | 0.2575 | 0.254 | 0.248 | 0.269 | 0.252 | 0.26074 | 0.256 | 0.2637 | 0.2714 | 0.22118 | > 0.2 | 0.2641 | 0.238 | 0.2518 | 0.2407 | 0.2841 | 0.2336 | 0.2609 | 0.228 | 0.244 |  | 1
| 1 | 2 | 0.26392 | 0.2524 | 0.251 | 0.242 | 0.269 | 0.256 | 0.2612 | 0.259 | 0.2612 | 0.2765 | 0.21379 | > 0.2 | 0.263 | 0.243 | 0.2379 | 0.2388 | 0.28863 | 0.242 | 0.2599 | 0.236 | 0.24 |  | 0
| 1 | 3 | 0.26474 | 0.2526 | 0.257 | 0.243 | 0.268 | 0.258 | 0.26262 | 0.263 | 0.2635 | 0.2777 | 0.21391 | > 0.2 | 0.2636 | 0.243 | 0.2444 | 0.2371 | 0.29791 | 0.239 | 0.2653 | 0.229 | 0.244 |  | 16
| 1 | 4 | 0.26579 | 0.2616 | 0.254 | 0.246 | 0.266 | 0.256 | 0.26672 | 0.25 | 0.2636 | 0.2664 | 0.21338 | > 0.2 | 0.2652 | 0.24 | 0.2487 | 0.2358 | 0.28571 | 0.2397 | 0.253 | 0.237 | 0.244 |  | 0.2536
| 1 | 5 | 0.26283 | 0.2566 | 0.26 | 0.242 | 0.265 | 0.264 | 0.25864 | 0.254 | 0.2652 | 0.2688 | 0.21353 | > 0.2 | 0.2555 | 0.241 | 0.2544 | 0.2375 | 0.29354 | 0.2391 | 0.2596 | 0.234 | 0.244 |  | 125
| 1 | 6 | 0.26934 | 0.2522 | 0.253 | 0.253 | 0.267 | 0.259 | 0.26324 | 0.253 | 0.2652 | 0.2742 | 0.21354 | > 0.2 | 0.2594 | 0.239 | 0.2481 | 0.2357 | 0.28968 | 0.2437 | 0.2649 | 0.24 | 0.24
| Mean |  | 0.26538 | 0.25548 | 0.25483 | 0.24567 | 0.26733 | 0.2575 | 0.26219 | 0.25583 | 0.26373 | 0.2725 | 0.21489 |  | 0.2618 | 0.24067 | 0.24755 | 0.2376 | 0.28993 | 0.23952 | 0.2606 | 0.234 | 0.24267
| Median |  | 0.26519 | 0.2546 | 0.254 | 0.2445 | 0.2675 | 0.257 | 0.26191 | 0.255 | 0.26365 | 0.2728 | 0.21367 |  | 0.2633 | 0.2405 | 0.2484 | 0.2373 | 0.28916 | 0.2394 | 0.2604 | 0.235 | 0.244
| Std Dev. |  | 0.0022365 | 0.003777 | 0.0031885 | 0.0043205 | 0.001633 | 0.0039875 | 0.0027375 | 0.0046224 | 0.0014692 | 0.0044281 | 0.0030868 |  | 0.0036601 | 0.0020656 | 0.005825 | 0.0019058 | 0.0051045 | 0.0034394 | 0.0044703 | 0.0046904 | 0.0020656
| Rel.Std.Dev. |  | 0.8428% | 1.478% | 1.251% | 1.759% | 0.6108% | 1.549% | 1.044% | 1.807% | 0.5571% | 1.625% | 1.436% |  | 1.398% | 0.8583% | 2.353% | 0.8021% | 1.761% | 1.436% | 1.715% | 2.004% | 0.8512%
| PDM3 |  | 4.643% | 0.7414% | 0.4851% | -3.129% | 5.414% | 1.537% | 3.388% | 0.8795% | 3.995% | 7.451% | -15.27% |  | 3.232% | -5.101% | -2.387% | -6.31% | 14.32% | -5.554% | 2.759% | -7.73% | -4.312%
| Z-Score (Absolute) |  | 0.66 | 0.01 | 0.06 | 0.68 | 0.8 | 0.13 | 0.45 | 0.01 | 0.55 | 1.15 | 2.88 | NA | 0.42 | 1.03 | 0.55 | 1.24 | 2.34 | 1.1 | 0.34 | 1.48 | 0.89
| Analytical results for Zr in OREAS 604 (Certified Value 21.1 ppm) | [23 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 25 | [9 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 1
|  |  | 0.5g | 40g | 20g | 0.15g | 0.5g | 01g | 0.5g | 01g | 10g | 0.25g | 10g | 0.25g | 40g | [9 empty] | 1
| 1 | 1 | 26.58 | 20.2 | 24.93 | 18 | 23 | 20.9912 | 21.08 | 24.57 | 19.5 | 20.2 | 18 | 18.607 | 17 | [9 empty] | 1
| 1 | 2 | 26.3825 | 20.8 | 25.4392 | 17 | 24 | 20.1088 | 21.86 | 25.42 | 19.8 | 20.3 | 19 | 18.96 | 17 | [9 empty] | 0
| 1 | 3 | 26.18 | 20.6 | 25.3496 | 16 | 24 | 20.6273 | 21.69 | 25.57 | 20.3 | 21.4 | 18 | 19.295 | 17 | [9 empty] | 16
| 1 | 4 | 25.8525 | 21.4 | 25.8427 | 16 | 24 | 20.5991 | 21.6 | 25.73 | 20.1 | 19.3 | 19 | 19.076 | 18 | [9 empty] | 21.1335
| 1 | 5 | 25.4125 | 20.3 | 24.8713 | 16 | 23 | 20.9861 | 21.14 | 23.74 | 20.3 | 20.5 | 19 | 19.193 | 18 | [9 empty] | 126
| 1 | 6 | 25.11 | 21.9 | 26.0622 | 17 | 23 | 21.0535 | 21.85 | 24.02 | 19.6 | 19.8 | 18 | 18.337 | 19
| Mean |  | 25.9196 | 20.8667 | 25.4158 | 16.6667 | 23.5 | 20.7277 | 21.5367 | 24.8417 | 19.9333 | 20.25 | 18.5 | 18.9113 | 17.6667
| Median |  | 26.0162 | 20.7 | 25.3944 | 16.5 | 23.5 | 20.8067 | 21.645 | 24.995 | 19.95 | 20.25 | 18.5 | 19.018 | 17.5
| Std Dev. |  | 0.57189 | 0.66232 | 0.47697 | 0.8165 | 0.54772 | 0.36111 | 0.34529 | 0.85049 | 0.35024 | 0.7064 | 0.54772 | 0.36837 | 0.8165
| Rel.Std.Dev. |  | 2.206% | 3.174% | 1.877% | 4.899% | 2.331% | 1.742% | 1.603% | 3.424% | 1.757% | 3.488% | 2.961% | 1.948% | 4.622%
| PDM3 |  | 22.65% | -1.263% | 20.26% | -21.14% | 11.2% | -1.92% | 1.908% | 17.55% | -5.679% | -4.181% | -12.46% | -10.52% | -16.4%
| Z-Score (Absolute) |  | 1.57 | 0.04 | 1.42 | 1.23 | 0.84 | 0 | 0.24 | 1.25 | 0.24 | 0.14 | 0.67 | 0.55 | 0.93
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C103: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C121: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C157: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C175: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C193: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C211: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C229: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C247: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C283: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C301: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C319: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C337: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C355: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C373: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C409: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C427: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C445: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C463: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C481: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C499: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C535: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C553: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C571: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C589: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C661: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1039: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1057: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1129: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1147: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1165: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1183: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1201: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on D57: Craig Hamlyn: The INAA data is omitted from determination of the certified value for Au and is used solely for the calculation of Tolerance Limits and homogeneity evaluation of OREAS 600.

## Laser Ablation
| Analytical results for Ag in OREAS 604 (Indicative Value 478 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 481 | 1
| 1 | 2 | 474 | 1
| Mean |  | 477.5 | 16
| Median |  | 477.5 | 477.5
| Std Dev. |  | 4.9497 | 7
| Rel.Std.Dev. |  | 1.037%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 604 (Indicative Value 952 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 999 | 1
| 1 | 2 | 904 | 2
| Mean |  | 951.5 | 16
| Median |  | 951.5 | 951.5
| Std Dev. |  | 67.1751 | 8
| Rel.Std.Dev. |  | 7.06%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 604 (Indicative Value 8380 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 8570 | 1
| 1 | 2 | 8190 | 3
| Mean |  | 8380 | 16
| Median |  | 8380 | 8380
| Std Dev. |  | 268.7006 | 9
| Rel.Std.Dev. |  | 3.206%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 604 (Indicative Value 0.25 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | < 0.2 | 4
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.25
| Std Dev. |  |  | 10
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 604 (Indicative Value 27.4 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27.5 | 1
| 1 | 2 | 27.2 | 5
| Mean |  | 27.35 | 16
| Median |  | 27.35 | 27.35
| Std Dev. |  | 0.21213 | 11
| Rel.Std.Dev. |  | 0.7756%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 604 (Indicative Value 15.6 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.5 | 1
| 1 | 2 | 16.6 | 6
| Mean |  | 15.55 | 16
| Median |  | 15.55 | 15.55
| Std Dev. |  | 1.4849 | 12
| Rel.Std.Dev. |  | 9.549%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 604 (Indicative Value 43.6 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 44.1 | 1
| 1 | 2 | 43.1 | 7
| Mean |  | 43.6 | 16
| Median |  | 43.6 | 43.6
| Std Dev. |  | 0.70711 | 13
| Rel.Std.Dev. |  | 1.622%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 604 (Indicative Value 44.9 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46.2 | 1
| 1 | 2 | 43.6 | 8
| Mean |  | 44.9 | 16
| Median |  | 44.9 | 44.9
| Std Dev. |  | 1.8385 | 14
| Rel.Std.Dev. |  | 4.095%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 604 (Indicative Value 37.5 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 37 | 1
| 1 | 2 | 38 | 9
| Mean |  | 37.5 | 16
| Median |  | 37.5 | 37.5
| Std Dev. |  | 0.70711 | 15
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 604 (Indicative Value 3.63 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.69 | 1
| 1 | 2 | 3.57 | 10
| Mean |  | 3.63 | 16
| Median |  | 3.63 | 3.63
| Std Dev. |  | 0.084853 | 16
| Rel.Std.Dev. |  | 2.338%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 604 (Indicative Value 20850 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 21500 | 1
| 1 | 2 | 20200 | 11
| Mean |  | 20850 | 16
| Median |  | 20850 | 20850
| Std Dev. |  | 919.2388 | 17
| Rel.Std.Dev. |  | 4.409%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 604 (Indicative Value 1.64 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.77 | 1
| 1 | 2 | 1.51 | 12
| Mean |  | 1.64 | 16
| Median |  | 1.64 | 1.64
| Std Dev. |  | 0.18385 | 18
| Rel.Std.Dev. |  | 11.21%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 604 (Indicative Value 0.64 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.73 | 1
| 1 | 2 | 0.54 | 13
| Mean |  | 0.635 | 16
| Median |  | 0.635 | 0.635
| Std Dev. |  | 0.13435 | 19
| Rel.Std.Dev. |  | 21.16%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 604 (Indicative Value 0.56 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.54 | 1
| 1 | 2 | 0.58 | 14
| Mean |  | 0.56 | 16
| Median |  | 0.56 | 0.56
| Std Dev. |  | 0.028284 | 20
| Rel.Std.Dev. |  | 5.051%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 604 (Indicative Value 27.8 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28.2 | 1
| 1 | 2 | 27.4 | 15
| Mean |  | 27.8 | 16
| Median |  | 27.8 | 27.8
| Std Dev. |  | 0.56569 | 21
| Rel.Std.Dev. |  | 2.035%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 604 (Indicative Value 2.65 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.42 | 1
| 1 | 2 | 2.87 | 16
| Mean |  | 2.645 | 16
| Median |  | 2.645 | 2.645
| Std Dev. |  | 0.3182 | 22
| Rel.Std.Dev. |  | 12.03%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 604 (Indicative Value 4.34 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.35 | 1
| 1 | 2 | 4.32 | 17
| Mean |  | 4.335 | 16
| Median |  | 4.335 | 4.335
| Std Dev. |  | 0.021213 | 23
| Rel.Std.Dev. |  | 0.4893%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 604 (Indicative Value 0.3 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.3 | 1
| Mean |  | 0.295 | 16
| Median |  | 0.295 | 0.295
| Std Dev. |  | 0.0070711 | 7
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 604 (Indicative Value 3.45 ppm) |  |  | Indicative
In | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.55 | 1
| 1 | 2 | 3.35 | 2
| Mean |  | 3.45 | 16
| Median |  | 3.45 | 3.45
| Std Dev. |  | 0.14142 | 8
| Rel.Std.Dev. |  | 4.099%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 604 (Indicative Value 25 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25.5 | 1
| 1 | 2 | 24.4 | 3
| Mean |  | 24.95 | 16
| Median |  | 24.95 | 24.95
| Std Dev. |  | 0.77782 | 9
| Rel.Std.Dev. |  | 3.118%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 604 (Indicative Value 0.08 ppm) |  |  | Indicative
Lu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.08 | 4
| Mean |  | 0.08 | 16
| Median |  | 0.08 | 0.08
| Std Dev. |  | 0 | 10
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 604 (Indicative Value 0.025 wt.%) |  |  | Indicative
Mn | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0253 | 1
| 1 | 2 | 0.0242 | 5
| Mean |  | 0.02475 | 16
| Median |  | 0.02475 | 0.02475
| Std Dev. |  | 0.00077782 | 11
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 604 (Indicative Value 5.1 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 5.6 | 6
| Mean |  | 5.1 | 16
| Median |  | 5.1 | 5.1
| Std Dev. |  | 0.70711 | 12
| Rel.Std.Dev. |  | 13.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 604 (Indicative Value 9.34 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.22 | 1
| 1 | 2 | 9.45 | 7
| Mean |  | 9.335 | 16
| Median |  | 9.335 | 9.335
| Std Dev. |  | 0.16263 | 13
| Rel.Std.Dev. |  | 1.742%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 604 (Indicative Value 18.2 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.6 | 1
| 1 | 2 | 17.8 | 8
| Mean |  | 18.2 | 16
| Median |  | 18.2 | 18.2
| Std Dev. |  | 0.56569 | 14
| Rel.Std.Dev. |  | 3.108%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 604 (Indicative Value 612 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 636 | 1
| 1 | 2 | 588 | 9
| Mean |  | 612 | 16
| Median |  | 612 | 612
| Std Dev. |  | 33.9411 | 15
| Rel.Std.Dev. |  | 5.546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 604 (Indicative Value 1011 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1030 | 1
| 1 | 2 | 991 | 10
| Mean |  | 1010.5 | 16
| Median |  | 1010.5 | 1010.5
| Std Dev. |  | 27.5772 | 16
| Rel.Std.Dev. |  | 2.729%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 604 (Indicative Value 5.04 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.09 | 1
| 1 | 2 | 4.99 | 11
| Mean |  | 5.04 | 16
| Median |  | 5.04 | 5.04
| Std Dev. |  | 0.070711 | 17
| Rel.Std.Dev. |  | 1.403%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 604 (Indicative Value 53 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 55.8 | 1
| 1 | 2 | 50.8 | 12
| Mean |  | 53.3 | 16
| Median |  | 53.3 | 53.3
| Std Dev. |  | 3.5355 | 18
| Rel.Std.Dev. |  | 6.633%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 604 (Indicative Value 0.008 ppm) |  |  | Indicative
Re | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | < 0.01 | 13
| Mean |  | 0.01 | 16
| Median |  | 0.01 | 0.0075
| Std Dev. |  |  | 19
Rel.Std.Dev.
| PDM3 |  | 33.33%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 604 (Indicative Value 155 ppm) |  |  | Indicative
Sb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 156 | 1
| 1 | 2 | 154 | 14
| Mean |  | 155 | 16
| Median |  | 155 | 155
| Std Dev. |  | 1.4142 | 20
| Rel.Std.Dev. |  | 0.9124%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 604 (Indicative Value 5 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 5.2 | 15
| Mean |  | 5 | 16
| Median |  | 5 | 5
| Std Dev. |  | 0.28284 | 21
| Rel.Std.Dev. |  | 5.657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 604 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 16
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 22
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 604 (Indicative Value 3.72 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.62 | 1
| 1 | 2 | 3.82 | 17
| Mean |  | 3.72 | 16
| Median |  | 3.72 | 3.72
| Std Dev. |  | 0.14142 | 23
| Rel.Std.Dev. |  | 3.802%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 604 (Indicative Value 4.5 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.4 | 1
| 1 | 2 | 4.6 | 1
| Mean |  | 4.5 | 16
| Median |  | 4.5 | 4.5
| Std Dev. |  | 0.14142 | 7
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 604 (Indicative Value 480 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 475 | 1
| 1 | 2 | 485 | 2
| Mean |  | 480 | 16
| Median |  | 480 | 480
| Std Dev. |  | 7.0711 | 8
| Rel.Std.Dev. |  | 1.473%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 604 (Indicative Value 0.83 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.81 | 3
| Mean |  | 0.825 | 16
| Median |  | 0.825 | 0.825
| Std Dev. |  | 0.021213 | 9
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 604 (Indicative Value 0.35 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.32 | 4
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.042426 | 10
| Rel.Std.Dev. |  | 12.12%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 604 (Indicative Value 25.4 ppm) |  |  | Indicative
Te | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 24.8 | 5
| Mean |  | 25.4 | 16
| Median |  | 25.4 | 25.4
| Std Dev. |  | 0.84853 | 11
| Rel.Std.Dev. |  | 3.341%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 604 (Indicative Value 8.69 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.63 | 1
| 1 | 2 | 8.74 | 6
| Mean |  | 8.685 | 16
| Median |  | 8.685 | 8.685
| Std Dev. |  | 0.077782 | 12
| Rel.Std.Dev. |  | 0.8956%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 604 (Indicative Value 0.192 wt.%) |  |  | Indicative
Ti | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.196 | 1
| 1 | 2 | 0.187 | 7
| Mean |  | 0.1915 | 16
| Median |  | 0.1915 | 0.1915
| Std Dev. |  | 0.006364 | 13
| Rel.Std.Dev. |  | 3.323%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 604 (Indicative Value 7.6 ppm) |  |  | Indicative
Tl | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.6 | 1
| 1 | 2 | 7.6 | 8
| Mean |  | 7.6 | 16
| Median |  | 7.6 | 7.6
| Std Dev. |  | 0 | 14
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 604 (Indicative Value 0.11 ppm) |  |  | Indicative
Tm | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.11 | 1
| 1 | 2 | 0.11 | 9
| Mean |  | 0.11 | 16
| Median |  | 0.11 | 0.11
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 604 (Indicative Value 3.11 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.08 | 1
| 1 | 2 | 3.14 | 10
| Mean |  | 3.11 | 16
| Median |  | 3.11 | 3.11
| Std Dev. |  | 0.042426 | 16
| Rel.Std.Dev. |  | 1.364%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 604 (Indicative Value 36.4 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.1 | 1
| 1 | 2 | 37.6 | 11
| Mean |  | 36.35 | 16
| Median |  | 36.35 | 36.35
| Std Dev. |  | 1.7678 | 17
| Rel.Std.Dev. |  | 4.863%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 604 (Indicative Value 16.2 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.8 | 1
| 1 | 2 | 16.6 | 12
| Mean |  | 16.2 | 16
| Median |  | 16.2 | 16.2
| Std Dev. |  | 0.56569 | 18
| Rel.Std.Dev. |  | 3.492%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 604 (Indicative Value 8.18 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.16 | 1
| 1 | 2 | 8.2 | 13
| Mean |  | 8.18 | 16
| Median |  | 8.18 | 8.18
| Std Dev. |  | 0.028284 | 19
| Rel.Std.Dev. |  | 0.3458%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 604 (Indicative Value 0.62 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.64 | 1
| 1 | 2 | 0.6 | 14
| Mean |  | 0.62 | 16
| Median |  | 0.62 | 0.62
| Std Dev. |  | 0.028284 | 20
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 604 (Indicative Value 2485 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2550 | 1
| 1 | 2 | 2420 | 15
| Mean |  | 2485 | 16
| Median |  | 2485 | 2485
| Std Dev. |  | 91.9239 | 21
| Rel.Std.Dev. |  | 3.699%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 604 (Indicative Value 146 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 142 | 1
| 1 | 2 | 150 | 16
| Mean |  | 146 | 16
| Median |  | 146 | 146
| Std Dev. |  | 5.6569 | 22
| Rel.Std.Dev. |  | 3.875%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 604 (Indicative Value 11.37 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.37 | 1
| 1 | 2 | 11.37 | 19
| Mean |  | 11.37 | 16
| Median |  | 11.37 | 11.37
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 604 (Indicative Value 960 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 950 | 1
| 1 | 2 | 970 | 2
| Mean |  | 960 | 16
| Median |  | 960 | 960
| Std Dev. |  | 14.1421 | 26
| Rel.Std.Dev. |  | 1.473%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 604 (Indicative Value 8635 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 8630 | 1
| 1 | 2 | 8640 | 3
| Mean |  | 8635 | 16
| Median |  | 8635 | 8635
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 0.08189%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 604 (Indicative Value 1.03 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 1.03 | 22
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.0070711 | 28
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 604 (Indicative Value 50 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 8
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 604 (Indicative Value 30 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 9
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 604 (Indicative Value 21850 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 21800 | 1
| 1 | 2 | 21900 | 11
| Mean |  | 21850 | 16
| Median |  | 21850 | 21850
| Std Dev. |  | 70.7107 | 31
| Rel.Std.Dev. |  | 0.3236%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 604 (Indicative Value 4.43 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.42 | 1
| 1 | 2 | 4.43 | 19
| Mean |  | 4.425 | 16
| Median |  | 4.425 | 4.425
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.1598%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 604 (Indicative Value 1.63 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.62 | 1
| 1 | 2 | 1.63 | 20
| Mean |  | 1.625 | 16
| Median |  | 1.625 | 1.625
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 0.4351%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 604 (Indicative Value 0.39 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.39 | 1
| 1 | 2 | 0.39 | 21
| Mean |  | 0.39 | 16
| Median |  | 0.39 | 0.39
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 604 (Indicative Value 0.04 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.04 | 1
| 1 | 2 | 0.04 | 22
| Mean |  | 0.04 | 16
| Median |  | 0.04 | 0.04
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 604 (Indicative Value 1.14 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.13 | 1
| 1 | 2 | 1.14 | 23
| Mean |  | 1.135 | 16
| Median |  | 1.135 | 1.135
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.623%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 604 (Indicative Value 625 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 620 | 1
| 1 | 2 | 630 | 9
| Mean |  | 625 | 16
| Median |  | 625 | 625
| Std Dev. |  | 7.0711 | 30
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 604 (Indicative Value 0.137 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.137 | 1
| 1 | 2 | 0.136 | 25
| Mean |  | 0.1365 | 16
| Median |  | 0.1365 | 0.1365
| Std Dev. |  | 0.00070711 | 31
| Rel.Std.Dev. |  | 0.518%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 604 (Indicative Value 1045 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1040 | 1
| 1 | 2 | 1050 | 10
| Mean |  | 1045 | 16
| Median |  | 1045 | 1045
| Std Dev. |  | 7.0711 | 25
| Rel.Std.Dev. |  | 0.6767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 604 (Indicative Value 66.72 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 66.68 | 1
| 1 | 2 | 66.76 | 20
| Mean |  | 66.72 | 16
| Median |  | 66.72 | 66.72
| Std Dev. |  | 0.056569 | 26
| Rel.Std.Dev. |  | 0.08478%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 604 (Indicative Value 12.5 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | < 10 | 1
| Mean |  | 20 | 16
| Median |  | 20 | 12.5
| Std Dev. |  |  | 27
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 604 (Indicative Value 11.45 wt.%) |  |  | Indicative
SO3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.45 | 1
| 1 | 2 | 11.45 | 22
| Mean |  | 11.45 | 16
| Median |  | 11.45 | 11.45
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 604 (Indicative Value 0.328 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.327 | 1
| 1 | 2 | 0.328 | 23
| Mean |  | 0.3275 | 16
| Median |  | 0.3275 | 0.3275
| Std Dev. |  | 0.00070711 | 29
| Rel.Std.Dev. |  | 0.2159%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 604 (Indicative Value < 10 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 10
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 30
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 604 (Indicative Value 2435 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 2430 | 1
| 1 | 2 | 2440 | 15
| Mean |  | 2435 | 16
| Median |  | 2435 | 2435
| Std Dev. |  | 7.0711 | 31
| Rel.Std.Dev. |  | 0.2904%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 604 (Indicative Value 7.3 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | at 1000°C | 2
2
| 1 | 1 | 7.28 | 1
| 1 | 2 | 7.31 | 27
| Mean |  | 7.295 | 16
| Median |  | 7.295 | 7.295
| Std Dev. |  | 0.021213 | 33
| Rel.Std.Dev. |  | 0.2908%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 604 (Certified Value 491 ppm) | [23 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | ppm
|  |  | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*AAS | 4A*OES/MS | 4A*AAS | 4A*AAS | 4A*OES | 4A*MS | 4A*AAS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 498.5867 | 489 | 471 | 500 | 505 | 487 | > 100 | 510.3 | 481 | > 100 | 494.7047 | 481.9 | 484 | N.A. | - | 518 | --- | 477.6989 | 494 | 486 | 485 | 411 | 1
| 1 | 2 | 502.0425 | 489 | 476 | 502 | 500 | 476 | > 100 | 494.094 | 481 | > 100 | 491.0063 | 478.3 | 485 | N.A. | - | 520 | --- | 483.12 | 488 | 491 | 495 | 427 | 0
| 1 | 3 | 502.308 | 476 | 479 | 506 | 504 | 493 | > 100 | 506.832 | 482 | > 100 | 505.7154 | 475.4 | 483.1 | N.A. | - | 523 | --- | 490.9872 |  | 497 | 480 | 408 | 16
| 1 | 4 | 498.0875 | 480 | 481 | 503 | 509 | 494 | > 100 | 499.857 | 497 | > 100 | 490.3376 | 493.7 | 479.9 | N.A. | - | 520 | --- | 466.7722 |  | 490 | 480 | 421 | 490.6509
| 1 | 5 | 496.77 | 489 | 475 | 498 | 502 | 487 | > 100 | 509.865 | 495 | > 100 | 494.5546 | 506 | 480.6 | N.A. | - | 528 | --- | 494.1888 |  | 497 | 500 | 410 | 10
| 1 | 6 | 497.51 | 485 | 493 | 499 | 499 | 479 | > 100 | 511.694 | 474 | > 100 | 489.2253 | 480.8 | 481.4 | N.A. | - | 534 | --- | 496.2048 |  | 487 | 500 | 414
| Mean |  | 499.2174 | 484.6667 | 479.1667 | 501.3333 | 503.1667 | 486 |  | 505.4403 | 485 |  | 494.2573 | 486.0167 | 482.3333 |  |  | 523.8333 |  | 484.8286 | 491 | 491.3333 | 490 | 415.1667
| Median |  | 498.3371 | 487 | 477.5 | 501 | 503 | 487 |  | 508.3485 | 481.5 |  | 492.7804 | 481.35 | 482.25 |  |  | 521.5 |  | 487.0536 | 491 | 490.5 | 490 | 412.5
| Std Dev. |  | 2.3712 | 5.5377 | 7.6004 | 2.9439 | 3.656 | 7.2664 |  | 6.9879 | 9.0111 |  | 6.047 | 11.6171 | 2.0136 |  |  | 6.08 |  | 11.2672 | 4.2426 | 4.761 | 9.4868 | 7.3598
| Rel.Std.Dev. |  | 0.475% | 1.143% | 1.586% | 0.5872% | 0.7266% | 1.495% |  | 1.383% | 1.858% |  | 1.223% | 2.39% | 0.4175% |  |  | 1.161% |  | 2.324% | 0.8641% | 0.969% | 1.936% | 1.773%
| PDM3 |  | 1.746% | -1.22% | -2.341% | 2.177% | 2.551% | -0.9479% |  | 3.014% | -1.152% |  | 0.735% | -0.9445% | -1.695% |  |  | 6.763% |  | -1.187% | 0.07115% | 0.1391% | -0.1327% | -15.38%
| Z-Score (Absolute) |  | 0.81 | 0.47 | 0.95 | 1 | 1.16 | 0.35 | NA | 1.36 | 0.44 | NA | 0.37 | 0.7 | 0.67 | NA | NA | 2.98 | NA | 0.45 | 0.09 | 0.12 | 0 | 6.58
| Analytical results for Al in OREAS 604 (Certified Value 5.82 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 5.665 | 5.6591 | 5.77 | 5.4 | 5.72 | 5.57 | 6.3209 | 6.64 | 6.02 | 3.52 | 5.9505 | 5.7594 | 6.14 | 5.35 | 5.9 | 5.15 | 5.853 | 3.84 | 5.8 | 5.8 |  |  | 1
| 1 | 2 | 5.715 | 5.7498 | 5.79 | 5.59 | 5.66 | 5.71 | 6.2703 | 7.47 | 6.01 | 3.47 | 5.9873 | 5.8162 | 6.16 | 5.24 | 5.88 | 5.52 | 5.907 | 4.93 | 5.68 | 5.79 |  |  | 0
| 1 | 3 | 5.715 | 5.7866 | 5.68 | 5.59 | 5.71 | 5.26 | 6.2489 | 7.21 | 5.99 | 3 | 6.0319 | 5.7835 | 6.18 | 4.95 | 6 | 5.42 | 5.833 | 4.56 | 5.67 | 5.95 |  |  | 16
| 1 | 4 | 5.725 | 5.7909 | 5.75 | 5.54 | 5.63 | 5.69 | 6.257 | 7.32 | 6 | 3.47 | 6.0465 | 5.8033 | 6.14 | 4.46 | 5.74 | 5.49 | 5.965 | 4.67 | 5.85 | 5.95 |  |  | 5.8168
| 1 | 5 | 5.625 | 5.74 | 5.63 | 5.8 | 5.51 | 5.67 | 6.3379 | 6.88 | 5.97 | 3.09 | 5.9613 | 5.8504 | 6.1 | 4.77 | 5.84 | 5.14 | 5.783 | 4.74 | 5.94 | 5.86 |  |  | 11
| 1 | 6 | 5.73 | 5.7504 | 5.68 | 5.51 | 5.8 | 5.71 | 6.3322 | 7.33 | 5.97 | 3.21 | 5.9693 | 5.8471 | 6.15 | 5.1 | 5.92 | 5.21 | 5.892 | 3.66 | 5.95 | 5.89
| Mean |  | 5.6958 | 5.7461 | 5.7167 | 5.5717 | 5.6717 | 5.6017 | 6.2945 | 7.1417 | 5.9933 | 3.2933 | 5.9911 | 5.81 | 6.145 | 4.9783 | 5.88 | 5.3217 | 5.8722 | 4.4 | 5.815 | 5.8733
| Median |  | 5.715 | 5.7501 | 5.715 | 5.565 | 5.685 | 5.68 | 6.2956 | 7.265 | 5.995 | 3.34 | 5.9783 | 5.8098 | 6.145 | 5.025 | 5.89 | 5.315 | 5.8725 | 4.615 | 5.825 | 5.875
| Std Dev. |  | 0.041763 | 0.047509 | 0.062503 | 0.13197 | 0.098268 | 0.17532 | 0.040167 | 0.31619 | 0.020656 | 0.22277 | 0.039376 | 0.03565 | 0.026646 | 0.32689 | 0.086718 | 0.17452 | 0.063367 | 0.52081 | 0.12211 | 0.070048
| Rel.Std.Dev. |  | 0.7332% | 0.8268% | 1.093% | 2.369% | 1.733% | 3.13% | 0.6381% | 4.427% | 0.3446% | 6.764% | 0.6572% | 0.6136% | 0.4336% | 6.566% | 1.475% | 3.279% | 1.079% | 11.84% | 2.1% | 1.193%
| PDM3 |  | -2.079% | -1.214% | -1.721% | -4.214% | -2.494% | -3.698% | 8.214% | 22.78% | 3.036% | -43.38% | 2.998% | -0.1165% | 5.643% | -14.41% | 1.087% | -8.511% | 0.9526% | -24.36% | -0.03021% | 0.9726%
| Z-Score (Absolute) |  | 0.35 | 0.14 | 0.26 | 0.89 | 0.46 | 0.46 | 2.22 | 5.85 | 0.92 | 10.66 | 0.91 | 0.14 | 1.57 | 3.43 | 0.44 | 1.96 | 0.4 | 5.91 | 0.16 | 0.41
| Analytical results for As in OREAS 604 (Certified Value 972 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 1036.8225 | 939.2 | 1040 | 992 | 1015 | 976 | 1044.72 | 980 | 1044 | 135.6 | 1067 | 985 | 914 | 972 | 887 | 921.059 | 874 | 835 | 912 | 946 |  |  | 1
| 1 | 2 | 1040.522 | 932.5 | 1030 | 1015 | 997 | 977 | 1044.4309 | 980 | 1025 | 123.2 | 1049 | 1052 | 902 | 927 | 930 | 950.318 | 878 | 876 | 896 | 983 |  |  | 0
| 1 | 3 | 1060.882 | 949.7 | 1010 | 990 | 999 | 1010 | 1025.8927 | 1000 | 1048 | 117 | 1036 | 953 | 876 | 981 | 959 | 930.289 | 894 | 879 | 904 | 964 |  |  | 16
| 1 | 4 | 1051.272 | 924.2 | 1010 | 987 | 998 | 1010 | 1033.545 | 1000 | 1060 | 126.8 | 1057 | 982 | 890 | 963 | 910 | 921.44 | 854 | 895 | 901 | 967 |  |  | 971.931
| 1 | 5 | 1054.156 | 940.8 | 997 | 1055 | 972 | 1030 | 1031.4854 | 970 | 1049 | 110.7 | 958 | 1011 | 916 | 1044 | 861 | 933.769 | 892 | 841 | 896 | 961 |  |  | 12
| 1 | 6 | 1066.376 | 936.8 | 997 | 983 | 1015 | 1040 | 1026.6664 | 980 | 1038 | 115.1 | 1041 | 1029 | 910 | 1080 | 883 | 934.091 | 903 | 844 | 944 | 998
| Mean |  | 1051.6717 | 937.2 | 1014 | 1003.6667 | 999.3333 | 1007.1667 | 1034.4567 | 985 | 1044 | 121.4 | 1034.6667 | 1002 | 901.3333 | 994.5 | 905 | 931.8277 | 882.5 | 861.6667 | 908.8333 | 969.8333
| Median |  | 1052.714 | 938 | 1010 | 991 | 998.5 | 1010 | 1032.5152 | 980 | 1046 | 120.1 | 1045 | 998 | 906 | 976.5 | 898.5 | 932.029 | 885 | 860 | 902.5 | 965.5
| Std Dev. |  | 11.421 | 8.5377 | 17.5613 | 27.5366 | 15.7818 | 26.4455 | 8.3488 | 12.2474 | 11.7813 | 9.0286 | 39.1646 | 35.7771 | 15.629 | 56.5641 | 35.5246 | 10.7426 | 17.5926 | 24.7683 | 18.2254 | 18.1925
| Rel.Std.Dev. |  | 1.086% | 0.911% | 1.732% | 2.744% | 1.579% | 2.626% | 0.8071% | 1.243% | 1.128% | 7.437% | 3.785% | 3.571% | 1.734% | 5.688% | 3.925% | 1.153% | 1.993% | 2.874% | 2.005% | 1.876%
| PDM3 |  | 8.204% | -3.573% | 4.328% | 3.265% | 2.819% | 3.625% | 6.433% | 1.345% | 7.415% | -87.51% | 6.455% | 3.094% | -7.264% | 2.322% | -6.886% | -4.126% | -9.201% | -11.34% | -6.492% | -0.2158%
| Z-Score (Absolute) |  | 0.79 | 0.66 | 0.31 | 0.05 | 0.13 | 0.23 | 0.57 | 0.05 | 0.69 | 10.96 | 0.77 | 0.16 | 1.11 | 0.07 | 1.06 | 0.72 | 1.35 | 1.61 | 1.1 | 0.24
| Analytical results for B in OREAS 604 (Indicative Value 13.3 ppm) | [23 empty] | Indicative
B | Round | Replicate | 01: Lab |  |  | 1
| No | No | 01: 25 |  |  | ppm
|  |  | 01: PF*OES |  |  | 1
1
| 1 | 1 | 01: 20 |  |  | 1
| 1 | 2 | 01: < 20 |  |  | 3
| 1 | 3 | 01: < 20 |  |  | 16
| 1 | 4 | 01: < 20 |  |  | 13.3333
| 1 | 5 | 01: < 20 |  |  | 9
| 1 | 6 | 01: 20
| Mean |  | 01: 20
| Median |  | 01: 20
| Std Dev. |  | 01: 0
Rel.Std.Dev. | bar: 01
| PDM3 |  | 01: 50%
Z-Score (Absolute) | bar: 01
| Analytical results for Ba in OREAS 604 (Indicative Value 5923 ppm) | [23 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 16 | 17 | 18 | 21 | 23 | 25 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*MS | [6 empty] | 0
0
| 1 | 1 | 3388.5 | 1410 | 50 | 160 | 2790 | 5667.3204 | 780 | 1252 | 1269.7494 | 1158 | 128 | 667 | 379 | 880 | 39 | 8560 | [6 empty] | 1
| 1 | 2 | 3386.3 | 1210 | 30 | 170 | 3500 | 5923.4784 | 866 | 1232 | 1297.5179 | 845 | 133 | 676 | 482 | 1037 | 26 | 8530 | [6 empty] | 4
| 1 | 3 | 3563.3 | 1160 | 30 | 170 | 110 | 5597.5846 | 848 | 1018 | 1148.7291 | 1030 | 123 | 693 | 387 | 1092 | 34 | 8500 | [6 empty] | 16
| 1 | 4 | 3604.4 | 1310 | 40 | 130 | 230 | 5762.2152 | 863 | 1140 | 1242.2547 | 1009 | 116 | 676 | 524 | 1292 | 32 | 8450 | [6 empty] | 5923.2928
| 1 | 5 | 3473.3 | 1690 | 40 | 140 | 220 | 5769.2959 | 870 | 898 | 1308.9335 | 2043 | 111 | 658 | 295 | 985 | 25 | 8590 | [6 empty] | 10
| 1 | 6 | 3576.1 | 1070 | 50 | 180 | 200 | 5707.475 | 844 | 991 | 1217.2953 | 1775 | 103 | 684 | 244 | 1118 | 42 | 8570
| Mean |  | 3498.65 | 1308.3333 | 40 | 158.3333 | 1175 | 5737.8949 | 845.1667 | 1088.5 | 1247.4133 | 1310 | 119 | 675.6667 | 385.1667 | 1067.3333 | 33 | 8533.3333
| Median |  | 3518.3 | 1260 | 40 | 165 | 225 | 5734.8451 | 855.5 | 1079 | 1256.002 | 1094 | 119.5 | 676 | 383 | 1064.5 | 33 | 8545
| Std Dev. |  | 96.7051 | 221.1259 | 8.9443 | 19.4079 | 1542.9679 | 111.0062 | 33.5405 | 141.9461 | 59.1039 | 482.0465 | 11.1535 | 12.3072 | 106.5372 | 138.8836 | 6.8118 | 51.6398
| Rel.Std.Dev. |  | 2.764% | 16.9% | 22.36% | 12.26% | 131.3% | 1.935% | 3.969% | 13.04% | 4.738% | 36.8% | 9.373% | 1.821% | 27.66% | 13.01% | 20.64% | 0.6052%
| PDM3 |  | -40.93% | -77.91% | -99.32% | -97.33% | -80.16% | -3.13% | -85.73% | -81.62% | -78.94% | -77.88% | -97.99% | -88.59% | -93.5% | -81.98% | -99.44% | 44.06%
| Z-Score (Absolute) |  | 3.7 | 0.5 | 1.35 | 1.17 | 0.37 | 6.96 | 0.15 | 0.18 | 0.42 | 0.51 | 1.23 | 0.42 | 0.84 | 0.15 | 1.36 | 11.04
| Analytical results for Be in OREAS 604 (Certified Value 1.13 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.1808 | 1.17 | 1.2 | 1.1 | 1.1 | 1.2 | 1.2264 | 1 | 1.35 | 1.0717 | 1.1 | 1.42 | 1.1 | 2.1 | < 3 | 0.983 | 2 | 1 | 1 | 1.1 |  |  | 1
| 1 | 2 | 1.2158 | 1.2 | 1.2 | 1.2 | 1.2 | 1.2 | 1.189 | 1.2 | 1.21 | 0.9914 | 1.1 | 1.45 | 1.1 | 2 | < 3 | 0.829 | 2 | 1 | 1 | 1.1 |  |  | 0
| 1 | 3 | 1.2199 | 1.17 | 1.1 | 1.2 | 1.2 | 1.1 | 1.1266 | 1.2 | 1.28 | 1.0061 | 1 | 1.26 | 1 | 2 | < 3 | 0.915 | 2 | 1 | 1 | 1.1 |  |  | 16
| 1 | 4 | 1.2184 | 1.16 | 1.1 | 1.2 | 1.1 | 1.1 | 1.1947 | 1.2 | 1.2 | 0.96707 | 1.1 | 1.32 | 1.1 | 2 | < 3 | 0.898 | 2 | 1 | 1 | 1.1 |  |  | 1.125
| 1 | 5 | 1.2197 | 1.14 | 1.1 | 1.2 | 1.1 | 1.1 | 1.161 | 1.3 | 1.2 | 1.0389 | 1.2 | 1.23 | 1.1 | 2.1 | < 3 | 0.95 | 2 | 1 | 1 | 1.1 |  |  | 13
| 1 | 6 | 1.2341 | 1.13 | 1.1 | 1.2 | 1.2 | 1.1 | 1.2137 | 1 | 1.36 | 1.0389 | 1 | 1.22 | 1.1 | 1.9 | < 3 | 1.152 | 2 | 1 | 1 | 1
| Mean |  | 1.2148 | 1.1617 | 1.1333 | 1.1833 | 1.15 | 1.1333 | 1.1852 | 1.15 | 1.2667 | 1.019 | 1.0833 | 1.3167 | 1.0833 | 2.0167 |  | 0.9545 | 2 | 1 | 1 | 1.0833
| Median |  | 1.219 | 1.165 | 1.1 | 1.2 | 1.15 | 1.1 | 1.1918 | 1.2 | 1.245 | 1.0225 | 1.1 | 1.29 | 1.1 | 2 |  | 0.9325 | 2 | 1 | 1 | 1.1
| Std Dev. |  | 0.017846 | 0.024833 | 0.05164 | 0.040825 | 0.054772 | 0.05164 | 0.036438 | 0.12247 | 0.074744 | 0.037933 | 0.075277 | 0.098522 | 0.040825 | 0.075277 |  | 0.10986 | 0 | 0 | 0 | 0.040825
| Rel.Std.Dev. |  | 1.469% | 2.138% | 4.556% | 3.45% | 4.763% | 4.556% | 3.074% | 10.65% | 5.901% | 3.723% | 6.949% | 7.483% | 3.768% | 3.733% |  | 11.51% | 0% | 0% | 0% | 3.768%
| PDM3 |  | 7.979% | 3.257% | 0.739% | 5.183% | 2.22% | 0.739% | 5.353% | 2.22% | 12.59% | -9.422% | -3.705% | 17.03% | -3.705% | 79.26% |  | -15.16% | 77.77% | -11.11% | -11.11% | -3.705%
| Z-Score (Absolute) |  | 0.72 | 0.12 | 0.17 | 0.34 | 0 | 0.17 | 0.36 | 0 | 1.18 | 1.32 | 0.67 | 1.69 | 0.67 | 8.77 | NA | 1.98 | 8.6 | 1.52 | 1.52 | 0.67
| Analytical results for Bi in OREAS 604 (Certified Value 28.4 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  | 1
1
| 1 | 1 | 38.62 | 30.66 | 36 | 26 | 35 | 33 | 27.1276 | 30 | 23 | 30.33 | 23.645 | 27.18 | 30.77 | < 30 | 28.5 | 45 | 26.065 | 26 | 26 | 25.1 | 27.9 |  | 1
| 1 | 2 | 35.7833 | 29.73 | 38 | 26 | 30 | 26 | 26.7407 | 50 | 24 | 27.39 | 24.9511 | 27.41 | 28.86 | < 30 | 25.8 | 48 | 25.495 | 28 | 24 | 25.5 | 28.6 |  | 0
| 1 | 3 | 36.2733 | 28.97 | 31 | 24 | 35 | 35 | 26.8661 | 50 | 25 | 28.83 | 25.098 | 26.82 | 29.11 | < 30 | 26.8 | 44 | 25.31 | 24 | 24 | 25 | 27.5 |  | 16
| 1 | 4 | 32.2167 | 30.05 | 36 | 27 | 31 | 36 | 26.4326 | 30 | 26 | 27.8 | 24.3179 | 27.01 | 27.48 | < 30 | 25.6 | 46 | 25.673 | 24 | 21 | 25.5 | 27.6 |  | 28.3856
| 1 | 5 | 35.86 | 29.71 | 37 | 33 | 30 | 35 | 27.1161 | 40 | 24 | 28.71 | 24.6648 | 27.02 | 29.38 | < 30 | 27.7 | 50 | 24.869 | 25 | 20 | 26 | 27.8 |  | 14
| 1 | 6 | 36.42 | 29.59 | 35 | 28 | 36 | 38 | 26.6359 | 30 | 24 | 29.86 | 24.2776 | 27 | 28.5 | < 30 | 26.1 | 45 | 25.741 | 20 | 24 | 25.4 | 27.5
| Mean |  | 35.8622 | 29.785 | 35.5 | 27.3333 | 32.8333 | 33.8333 | 26.8198 | 38.3333 | 24.3333 | 28.82 | 24.4924 | 27.0733 | 29.0167 |  | 26.75 | 46.3333 | 25.5255 | 24.5 | 23.1667 | 25.4167 | 27.8167
| Median |  | 36.0667 | 29.72 | 36 | 26.5 | 33 | 35 | 26.8034 | 35 | 24 | 28.77 | 24.4914 | 27.015 | 28.985 |  | 26.45 | 45.5 | 25.584 | 24.5 | 24 | 25.45 | 27.7
| Std Dev. |  | 2.0679 | 0.55619 | 2.429 | 3.0768 | 2.7869 | 4.1673 | 0.27373 | 9.8319 | 1.0328 | 1.1366 | 0.52943 | 0.20057 | 1.0833 |  | 1.1502 | 2.2509 | 0.4093 | 2.6646 | 2.2286 | 0.35449 | 0.41673
| Rel.Std.Dev. |  | 5.766% | 1.867% | 6.842% | 11.26% | 8.488% | 12.32% | 1.021% | 25.65% | 4.244% | 3.944% | 2.162% | 0.7408% | 3.733% |  | 4.3% | 4.858% | 1.603% | 10.88% | 9.62% | 1.395% | 1.498%
| PDM3 |  | 26.34% | 4.93% | 25.06% | -3.707% | 15.67% | 19.19% | -5.516% | 35.05% | -14.28% | 1.53% | -13.72% | -4.623% | 2.223% |  | -5.762% | 63.23% | -10.08% | -13.69% | -18.39% | -10.46% | -2.004%
| Z-Score (Absolute) |  | 2.33 | 0.6 | 2.29 | 0.32 | 1.37 | 2.03 | 0.16 | 2.78 | 0.79 | 0.35 | 0.75 | 0.09 | 0.4 | NA | 0.18 | 4.82 | 0.49 | 0.75 | 1.09 | 0.52 | 0.09
| Analytical results for Ca in OREAS 604 (Certified Value 0.735 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.79 | 0.691 | 0.75 | 0.72 | 0.76 | 0.74 | 0.75267 | 0.8 | 0.75 | 0.7 | 0.69088 | 0.7322 | 0.75 | 0.69 | 0.71 | 0.64 | 0.729 | 0.46 | 0.74 | 0.74 |  |  | 1
| 1 | 2 | 0.795 | 0.6962 | 0.76 | 0.75 | 0.76 | 0.74 | 0.73777 | 0.86 | 0.74 | 0.64 | 0.70479 | 0.7329 | 0.76 | 0.68 | 0.73 | 0.68 | 0.725 | 0.6 | 0.75 | 0.75 |  |  | 0
| 1 | 3 | 0.775 | 0.6976 | 0.74 | 0.73 | 0.77 | 0.69 | 0.75436 | 0.85 | 0.75 | 0.68 | 0.69427 | 0.736 | 0.75 | 0.65 | 0.75 | 0.68 | 0.699 | 0.56 | 0.72 | 0.75 |  |  | 16
| 1 | 4 | 0.745 | 0.6987 | 0.74 | 0.73 | 0.75 | 0.74 | 0.73718 | 0.86 | 0.75 | 0.7 | 0.70808 | 0.7389 | 0.75 | 0.6 | 0.71 | 0.68 | 0.736 | 0.58 | 0.75 | 0.74 |  |  | 0.73504
| 1 | 5 | 0.735 | 0.7218 | 0.72 | 0.77 | 0.73 | 0.74 | 0.75352 | 0.81 | 0.76 | 0.7 | 0.68911 | 0.7446 | 0.76 | 0.65 | 0.72 | 0.6 | 0.71 | 0.59 | 0.75 | 0.75 |  |  | 15
| 1 | 6 | 0.79 | 0.7089 | 0.73 | 0.73 | 0.77 | 0.75 | 0.75544 | 0.88 | 0.72 | 0.72 | 0.69861 | 0.7342 | 0.77 | 0.66 | 0.73 | 0.65 | 0.701 | 0.43 | 0.76 | 0.74
| Mean |  | 0.77167 | 0.70237 | 0.74 | 0.73833 | 0.75667 | 0.73333 | 0.74849 | 0.84333 | 0.745 | 0.69 | 0.69762 | 0.73647 | 0.75667 | 0.655 | 0.725 | 0.655 | 0.71667 | 0.53667 | 0.745 | 0.745
| Median |  | 0.7825 | 0.69815 | 0.74 | 0.73 | 0.76 | 0.74 | 0.7531 | 0.855 | 0.75 | 0.7 | 0.69644 | 0.7351 | 0.755 | 0.655 | 0.725 | 0.665 | 0.7175 | 0.57 | 0.75 | 0.745
| Std Dev. |  | 0.025626 | 0.011169 | 0.014142 | 0.018348 | 0.015055 | 0.021602 | 0.0085854 | 0.031411 | 0.013784 | 0.027568 | 0.0076261 | 0.0046552 | 0.008165 | 0.031464 | 0.015166 | 0.032094 | 0.015475 | 0.072847 | 0.013784 | 0.0054772
| Rel.Std.Dev. |  | 3.321% | 1.59% | 1.911% | 2.485% | 1.99% | 2.946% | 1.147% | 3.725% | 1.85% | 3.995% | 1.093% | 0.6321% | 1.079% | 4.804% | 2.092% | 4.9% | 2.159% | 13.57% | 1.85% | 0.7352%
| PDM3 |  | 4.983% | -4.445% | 0.675% | 0.4483% | 2.942% | -0.232% | 1.83% | 14.73% | 1.355% | -6.127% | -5.09% | 0.1943% | 2.942% | -10.89% | -1.366% | -10.89% | -2.499% | -26.99% | 1.355% | 1.355%
| Z-Score (Absolute) |  | 1.22 | 1.31 | 0.06 | 0.23 | 0.67 | 0.14 | 0.38 | 3.84 | 0.25 | 1.4 | 1.49 | 0.06 | 0.67 | 3.04 | 0.48 | 3.04 | 0.79 | 7.37 | 0.25 | 0.25
| Analytical results for Cd in OREAS 604 (Certified Value 14.9 ppm) | [23 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 24.9201 | 14.54 | 15.6 | 15.5 | 15.9 | 15.8 | 15.3806 | 20 | 14 | 17.64 | 13.5919 | 15.75 | 16.55 | 15 | 15.7 | 26 | 13.78 | 13 | 14.6 | 13 | 15 |  | 1
| 1 | 2 | 25.6178 | 14.16 | 15.2 | 15.7 | 15.7 | 15.8 | 15.8082 | 24 | 15 | 16.36 | 14.1214 | 15.62 | 14.91 | 15 | 15.1 | 28 | 13.474 | 13 | 14.6 | 13 | 15 |  | 0
| 1 | 3 | 25.7414 | 14.09 | 15.6 | 15.5 | 15.7 | 15 | 15.411 | 17 | 14 | 16.95 | 14.4301 | 15.16 | 16.2 | 14 | 15.5 | 29 | 13.526 | 13 | 14.5 | 13 | 16 |  | 16
| 1 | 4 | 25.8167 | 14.22 | 15.4 | 15.6 | 15.7 | 15.2 | 15.8343 | 18 | 15 | 16.69 | 13.7741 | 15.75 | 15.08 | 14 | 15.1 | 27 | 13.792 | 13 | 15 | 13.5 | 15 |  | 14.8721
| 1 | 5 | 25.828 | 14.47 | 15.5 | 16.6 | 15 | 15 | 15.6913 | 20 | 15 | 16.91 | 13.573 | 15.46 | 15.95 | 15 | 15.7 | 28 | 13.352 | 13 | 14.5 | 13 | 15 |  | 16
| 1 | 6 | 26.0076 | 14.18 | 15.2 | 15.9 | 15.7 | 15.4 | 15.827 | 23 | 13 | 17.07 | 14.3924 | 15.45 | 15.09 | 15 | 15.4 | 27 | 13.34 | 13 | 14.4 | 13 | 15
| Mean |  | 25.6553 | 14.2767 | 15.4167 | 15.8 | 15.6167 | 15.3667 | 15.6588 | 20.3333 | 14.3333 | 16.9367 | 13.9805 | 15.5317 | 15.63 | 14.6667 | 15.4167 | 27.5 | 13.544 | 13 | 14.6 | 13.0833 | 15.1667
| Median |  | 25.7791 | 14.2 | 15.45 | 15.65 | 15.7 | 15.3 | 15.7498 | 20 | 14.5 | 16.93 | 13.9477 | 15.54 | 15.52 | 15 | 15.45 | 27.5 | 13.5 | 13 | 14.55 | 13 | 15
| Std Dev. |  | 0.38193 | 0.18316 | 0.18348 | 0.41952 | 0.31252 | 0.36697 | 0.21035 | 2.7325 | 0.8165 | 0.42585 | 0.38765 | 0.22498 | 0.69083 | 0.5164 | 0.27142 | 1.0488 | 0.20045 | 0 | 0.20976 | 0.20412 | 0.40825
| Rel.Std.Dev. |  | 1.489% | 1.283% | 1.19% | 2.655% | 2.001% | 2.388% | 1.343% | 13.44% | 5.696% | 2.514% | 2.773% | 1.449% | 4.42% | 3.521% | 1.761% | 3.814% | 1.48% | 0% | 1.437% | 1.56% | 2.692%
| PDM3 |  | 72.51% | -4.004% | 3.661% | 6.239% | 5.006% | 3.325% | 5.289% | 36.72% | -3.623% | 13.88% | -5.995% | 4.435% | 5.096% | -1.381% | 3.661% | 84.91% | -8.93% | -12.59% | -1.83% | -12.03% | 1.981%
| Z-Score (Absolute) |  | 8.45 | 0.94 | 0 | 0.18 | 0.17 | 0.04 | 0.2 | 4.06 | 0.89 | 1.26 | 1.19 | 0.09 | 0.18 | 0.62 | 0 | 9.98 | 1.55 | 2 | 0.67 | 1.93 | 0.34
| Analytical results for Ce in OREAS 604 (Certified Value 38.1 ppm) | [23 empty] | Certified
Ce | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 13; 08: 15; 09: 17; 10: 19; 11: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS |  | 1
1
| 1 | 1 | 01: 35.24; 02: 42.1703; 05: 36; 06: 20.16; 07: 43.9455; 08: 43.02; 09: 35.2; 10: 34.18; 11: 36 |  | 1
| 1 | 2 | 01: 38.13; 02: 40.3514; 05: 35; 06: 20.1; 07: 43.9039; 08: 44.22; 09: 34.1; 10: 32.99; 11: 34.7 |  | 0
| 1 | 3 | 01: 36.32; 02: 41.6479; 05: 36; 06: 17.18; 07: 42.9911; 08: 41.44; 09: 33.7; 10: 31.542; 11: 37.1 |  | 16
| 1 | 4 | 01: 36.11; 02: 40.5874; 05: 37; 06: 20.9; 07: 44.7862; 08: 42.66; 09: 37.5; 10: 33.493; 11: 35.7 |  | 38.0843
| 1 | 5 | 01: 36.25; 02: 41.7653; 05: 37; 06: 14.35; 07: 45.4919; 08: 42.06; 09: 41.3; 10: 33.029; 11: 37.5 |  | 17
| 1 | 6 | 01: 36.73; 02: 42.4843; 05: 35; 06: 21.72; 07: 44.4019; 08: 42.56; 09: 35; 10: 31.943; 11: 34
| Mean |  | 01: 36.4633; 02: 41.5011; 05: 36; 06: 19.0683; 07: 44.2534; 08: 42.66; 09: 36.1333; 10: 32.8628; 11: 35.8333
| Median |  | 01: 36.285; 02: 41.7066; 05: 36; 06: 20.13; 07: 44.1737; 08: 42.61; 09: 35.1; 10: 33.0095; 11: 35.85
| Std Dev. |  | 01: 0.95251; 02: 0.85586; 05: 0.89443; 06: 2.7738; 07: 0.85441; 08: 0.94038; 09: 2.8556; 10: 0.97637; 11: 1.3471
| Rel.Std.Dev. |  | 01: 2.612%; 02: 2.062%; 05: 2.485%; 06: 14.55%; 07: 1.931%; 08: 2.204%; 09: 7.903%; 10: 2.971%; 11: 3.759%
| PDM3 |  | 01: -4.256%; 02: 8.972%; 05: -5.473%; 06: -49.93%; 07: 16.2%; 08: 12.01%; 09: -5.123%; 10: -13.71%; 11: -5.91%
| Z-Score (Absolute) |  | 01: 0.1; 02: 1.18; 05: 0; 06: 3.44; 07: 1.77; 08: 1.43; 09: 0.19; 10: 0.67; 11: 0.04
| Analytical results for Co in OREAS 604 (Certified Value 42 ppm) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES |  | 1
2
| 1 | 1 | 42.5115 | 40.8 | 45 | 44 | 44 | 43 | 44.1803 | 41 | 44 | 15.5 | 43.6158 | 43.4 | 44 | 38 | 40.5 | 34 | 38.467 | 41 | 42 | 41 | 45 |  | 1
| 1 | 2 | 43.8036 | 41 | 45 | 47 | 45 | 43 | 43.9834 | 37 | 44 | 17 | 43.5001 | 42.7 | 44 | 38 | 38.7 | 37 | 37.717 | 40 | 42 | 40 | 45 |  | 0
| 1 | 3 | 43.9484 | 40.8 | 44 | 45 | 44 | 40 | 43.3146 | 39 | 45 | 13.1 | 43.605 | 42.2 | 45 | 36 | 40.6 | 38 | 38.187 | 40 | 43 | 40 | 45 |  | 16
| 1 | 4 | 44.1192 | 40.2 | 44 | 44 | 44 | 42 | 43.6239 | 36 | 45 | 16.2 | 43.7034 | 43.3 | 45 | 37 | 40.9 | 36 | 39.153 | 40 | 43 | 42 | 45 |  | 41.971
| 1 | 5 | 44.7194 | 40.8 | 43 | 48 | 42 | 42 | 44.0809 | 38 | 45 | 13.5 | 43.5039 | 44.8 | 44 | 38 | 44.4 | 35 | 38.297 | 41 | 41 | 40 | 45 |  | 18
| 1 | 6 | 44.6418 | 40.7 | 44 | 44 | 44 | 43 | 43.5052 | 39 | 43 | 13.7 | 43.6908 | 43 | 44 | 37 | 44.4 | 35 | 38.446 | 41 | 42 | 42 | 45
| Mean |  | 43.9573 | 40.7167 | 44.1667 | 45.3333 | 43.8333 | 42.1667 | 43.7814 | 38.3333 | 44.3333 | 14.8333 | 43.6032 | 43.2333 | 44.3333 | 37.3333 | 41.5833 | 35.8333 | 38.3778 | 40.5 | 42.1667 | 40.8333 | 45
| Median |  | 44.0338 | 40.8 | 44 | 44.5 | 44 | 42.5 | 43.8036 | 38.5 | 44.5 | 14.6 | 43.6104 | 43.15 | 44 | 37.5 | 40.75 | 35.5 | 38.3715 | 40.5 | 42 | 40.5 | 45
| Std Dev. |  | 0.7987 | 0.27142 | 0.75277 | 1.7512 | 0.98319 | 1.169 | 0.34888 | 1.7512 | 0.8165 | 1.617 | 0.087622 | 0.88242 | 0.5164 | 0.8165 | 2.3147 | 1.472 | 0.46758 | 0.54772 | 0.75277 | 0.98319 | 0
| Rel.Std.Dev. |  | 1.817% | 0.6666% | 1.704% | 3.863% | 2.243% | 2.772% | 0.7969% | 4.568% | 1.842% | 10.9% | 0.201% | 2.041% | 1.165% | 2.187% | 5.566% | 4.108% | 1.218% | 1.352% | 1.785% | 2.408% | 0%
| PDM3 |  | 4.733% | -2.989% | 5.231% | 8.011% | 4.437% | 0.4662% | 4.313% | -8.667% | 5.629% | -64.66% | 3.889% | 3.008% | 5.629% | -11.05% | -0.9236% | -14.62% | -8.561% | -3.505% | 0.4662% | -2.711% | 7.217%
| Z-Score (Absolute) |  | 0.67 | 0.55 | 0.75 | 1.19 | 0.63 | 0 | 0.61 | 1.44 | 0.82 | 10.29 | 0.54 | 0.4 | 0.82 | 1.82 | 0.22 | 2.38 | 1.43 | 0.63 | 0 | 0.5 | 1.07
| Analytical results for Cr in OREAS 604 (Certified Value 33.8 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES |  | 1
1
| 1 | 1 | 30.5 | 32 | 23 | 30 | 38 | 33 | 33.8 | 34 | 35 | 31 | 36.9853 | 36 | 35.93 | 32 | 20 | 13 | 34.106 | 39 | 30 | 40 | 50 |  | 1
| 1 | 2 | 32.7267 | 32 | 25 | 30 | 30 | 29 | 37.4483 | 33 | 37 | 25 | 38.5061 | 35 | 35.51 | 30 | 18 | 15 | 33.864 | 37 | 29 | 30 | 50 |  | 0
| 1 | 3 | 32.15 | 30 | 22 | 32 | 34 | 33 | 34.4224 | 36 | 37 | 30 | 40.9293 | 35 | 34.38 | 33 | 23 | 18 | 31.064 | 37 | 35 | 20 | < 50 |  | 16
| 1 | 4 | 31.02 | 32 | 24 | 28 | 34 | 31 | 35.435 | 34 | 35 | 28 | 37.356 | 36 | 36.04 | 36 | 24 | 12 | 26.131 | 37 | 45 | 30 | < 50 |  | 33.8189
| 1 | 5 | 30.4 | 31 | 31 | 31 | 35 | 30 | 34.6147 | 35 | 33 | 27 | 36.8804 | 39 | 36.13 | 40 | 20 | 16 | 34.293 | 42 | 34 | 30 | 50 |  | 19
| 1 | 6 | 33.3 | 30 | 25 | 31 | 37 | 30 | 36.9936 | 36 | 37 | 27 | 41.3956 | 40 | 34.73 | 49 | 16 | 16 | 34.456 | 43 | 34 | 30 | 50
| Mean |  | 31.6828 | 31.1667 | 25 | 30.3333 | 34.6667 | 31 | 35.4523 | 34.6667 | 35.6667 | 28 | 38.6754 | 36.8333 | 35.4533 | 36.6667 | 20.1667 | 15 | 32.319 | 39.1667 | 34.5 | 30 | 50
| Median |  | 31.585 | 31.5 | 24.5 | 30.5 | 34.5 | 30.5 | 35.0248 | 34.5 | 36 | 27.5 | 37.931 | 36 | 35.72 | 34.5 | 20 | 15.5 | 33.985 | 38 | 34 | 30 | 50
| Std Dev. |  | 1.2171 | 0.98319 | 3.1623 | 1.3663 | 2.8048 | 1.6733 | 1.4733 | 1.2111 | 1.633 | 2.1909 | 2.0163 | 2.137 | 0.7359 | 6.9761 | 2.9944 | 2.1909 | 3.2836 | 2.7142 | 5.6833 | 6.3246 | 0
| Rel.Std.Dev. |  | 3.842% | 3.155% | 12.65% | 4.504% | 8.091% | 5.398% | 4.156% | 3.493% | 4.578% | 7.825% | 5.213% | 5.802% | 2.076% | 19.03% | 14.85% | 14.61% | 10.16% | 6.93% | 16.47% | 21.08% | 0%
| PDM3 |  | -6.316% | -7.842% | -26.08% | -10.31% | 2.507% | -8.335% | 4.83% | 2.507% | 5.464% | -17.21% | 14.36% | 8.914% | 4.833% | 8.421% | -40.37% | -55.65% | -4.435% | 15.81% | 2.014% | -11.29% | 47.85%
| Z-Score (Absolute) |  | 0.53 | 0.67 | 2.75 | 0.91 | 0.31 | 0.72 | 0.53 | 0.31 | 0.6 | 1.57 | 1.44 | 0.92 | 0.54 | 0.18 | 3.78 | 5.24 | 0 | 1.58 | 0.33 | 0.44 | 4.64
| Analytical results for Cs in OREAS 604 (Certified Value 3.63 ppm) | [23 empty] | Certified
Cs | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 24; 12: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS |  | 2
3
| 1 | 1 | 01: 3.41; 02: 3.6706; 05: 4; 06: 3.5904; 07: 3.7; 08: 3.9; 09: 3.75; 10: 3.119; 11: 3.7; 12: 3.3 |  | 1
| 1 | 2 | 01: 3.41; 02: 3.5942; 05: 3.76; 06: 3.8317; 07: 3.6; 08: 3.51; 09: 3.41; 10: 3.109; 11: 3.6; 12: 3.4 |  | 0
| 1 | 3 | 01: 3.39; 02: 3.6158; 05: 3.98; 06: 3.6163; 07: 3.8; 08: 3.81; 09: 3.52; 10: 3.021; 11: 3.6; 12: 3.4 |  | 16
| 1 | 4 | 01: 3.51; 02: 3.4697; 05: 3.86; 06: 3.6641; 07: 3.7; 08: 3.47; 09: 3.56; 10: 3.15; 11: 3.8; 12: 3.4 |  | 3.6346
| 1 | 5 | 01: 3.53; 02: 3.5359; 05: 3.86; 06: 3.9484; 07: 3.6; 08: 3.8; 09: 3.64; 10: 3.102; 11: 3.7; 12: 3.4 |  | 20
| 1 | 6 | 01: 3.44; 02: 3.5614; 05: 4.14; 06: 3.7181; 07: 3.7; 08: 3.6; 09: 3.49; 10: 3.06; 11: 3.9; 12: 3.4
| Mean |  | 01: 3.4483; 02: 3.5746; 05: 3.9333; 06: 3.7282; 07: 3.6833; 08: 3.6817; 09: 3.5617; 10: 3.0935; 11: 3.7167; 12: 3.3833
| Median |  | 01: 3.425; 02: 3.5778; 05: 3.92; 06: 3.6911; 07: 3.7; 08: 3.7; 09: 3.54; 10: 3.1055; 11: 3.7; 12: 3.4
| Std Dev. |  | 01: 0.058109; 02: 0.069224; 05: 0.13426; 06: 0.13779; 07: 0.075277; 08: 0.17837; 09: 0.11957; 10: 0.045899; 11: 0.1169; 12: 0.040825
| Rel.Std.Dev. |  | 01: 1.685%; 02: 1.937%; 05: 3.413%; 06: 3.696%; 07: 2.044%; 08: 4.845%; 09: 3.357%; 10: 1.484%; 11: 3.145%; 12: 1.207%
| PDM3 |  | 01: -5.124%; 02: -1.65%; 05: 8.22%; 06: 2.575%; 07: 1.342%; 08: 1.296%; 09: -2.006%; 10: -14.89%; 11: 2.259%; 12: -6.912%
| Z-Score (Absolute) |  | 01: 1.29; 02: 0.38; 05: 2.18; 06: 0.72; 07: 0.39; 08: 0.38; 09: 0.48; 10: 3.82; 11: 0.63; 12: 1.75
| Analytical results for Cu in OREAS 604 (Certified Value 2.16 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*MS | PF*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.1301 | 2.1421 | 2.204 | 2.17 | 2.21 | 2.14 | 2.36 | 2.1355 | 2.13 | 2.268 | 2.12 | 2.1369 | 2.1529 | 2.1643 | 2.01 | > 10000 | > 10000 | 2.18 | 2.0101 | 2.092 | 2.14 | 2.18 | 1
| 1 | 2 | 2.1354 | 2.147 | 2.227 | 2.18 | 2.19 | 2.11 | 2.3 | 2.1296 | 2.14 | 2.262 | 2.23 | 2.1739 | 2.2036 | 2.1768 | 2.24 | > 10000 | -- | 2.14 | 2.0257 | 2.071 | 2.14 | 2.21 | 0
| 1 | 3 | 2.1228 | 2.1225 | 2.25 | 2.19 | 2.21 | 2.18 | 2.32 | 2.1439 | 2.2 | 2.329 | 2.15 | 2.1781 | 2.1675 | 2.0777 | 2.23 | > 10000 | > 10000 | 2.19 | 2.0956 | 2.075 | 2.18 | 2.18 | 16
| 1 | 4 | 2.131 | 2.1187 | 2.213 | 2.2 | 2.24 | 2.17 | 2.26 | 2.1109 | 2.2 | 2.335 | 2.16 | 2.1502 | 2.1383 | 2.1933 | 2.27 | > 10000 | > 10000 | 2.18 | 2.1221 | 2.092 | 2.17 | 2.19 | 2.1641
| 1 | 5 | 2.1257 | 2.1293 | 2.229 | 2.17 | 2.19 | 2.17 | 2.26 | 2.1607 | 2.18 | 2.325 | 2.28 | 2.1142 | 2.1791 | 2.045 | 2.28 | > 10000 | > 10000 | 2.177 | 2.1252 | 2.119 | 2.2 | 2.17 | 21
| 1 | 6 | 2.1271 | 2.1264 | 2.289 | 2.17 | 2.19 | 2.11 | 2.25 | 2.1348 | 2.17 | 2.295 | 2.21 | 2.132 | 2.1781 | 2.2043 | 2.14 | > 10000 | > 10000 | 2.14 | 2.2157 | 2.093 | 2.22 | 2.19
| Mean |  | 2.1287 | 2.131 | 2.2353 | 2.18 | 2.205 | 2.1467 | 2.2917 | 2.1359 | 2.17 | 2.3023 | 2.1917 | 2.1476 | 2.1699 | 2.1435 | 2.195 |  |  | 2.1678 | 2.0991 | 2.0903 | 2.175 | 2.1867
| Median |  | 2.1286 | 2.1279 | 2.228 | 2.175 | 2.2 | 2.155 | 2.28 | 2.1352 | 2.175 | 2.31 | 2.185 | 2.1435 | 2.1728 | 2.1705 | 2.235 |  |  | 2.1785 | 2.1088 | 2.092 | 2.175 | 2.185
| Std Dev. |  | 0.0044293 | 0.011196 | 0.030612 | 0.012649 | 0.019748 | 0.031411 | 0.043089 | 0.016408 | 0.029665 | 0.032098 | 0.059133 | 0.024926 | 0.02272 | 0.065951 | 0.1033 |  |  | 0.022004 | 0.075002 | 0.01699 | 0.032094 | 0.013663
| Rel.Std.Dev. |  | 0.2081% | 0.5254% | 1.369% | 0.5802% | 0.8956% | 1.463% | 1.88% | 0.7682% | 1.367% | 1.394% | 2.698% | 1.161% | 1.047% | 3.077% | 4.706% |  |  | 1.015% | 3.573% | 0.8128% | 1.476% | 0.6248%
| PDM3 |  | -1.637% | -1.53% | 3.291% | 0.7343% | 1.89% | -0.806% | 5.894% | -1.303% | 0.2722% | 6.387% | 1.273% | -0.765% | 0.2684% | -0.9503% | 1.427% |  |  | 0.1721% | -3.005% | -3.409% | 0.5033% | 1.042%
| Z-Score (Absolute) |  | 0.92 | 0.87 | 1.46 | 0.22 | 0.78 | 0.52 | 2.71 | 0.76 | 0 | 2.95 | 0.48 | 0.5 | 0 | 0.59 | 1.38 | NA | NA | 0.05 | 1.18 | 1.78 | 0.11 | 0.37
| Analytical results for Dy in OREAS 604 (Certified Value 1.6 ppm) | [23 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 02; 02: 09; 04: 13; 05: 15; 06: 17; 07: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS | 2
3
| 1 | 1 | 01: 1.59; 02: 1.702; 04: 1.4022; 05: 1.74; 06: 1.61; 07: 1.6 | 1
| 1 | 2 | 01: 1.59; 02: 1.6945; 04: 1.5267; 05: 1.68; 06: 1.54; 07: 1.5 | 0
| 1 | 3 | 01: 1.58; 02: 1.6745; 04: 1.3714; 05: 1.66; 06: 1.55; 07: 1.35 | 16
| 1 | 4 | 01: 1.6; 02: 1.7309; 04: 1.5458; 05: 1.69; 06: 1.53; 07: 1.55 | 1.6008
| 1 | 5 | 01: 1.66; 02: 1.7355; 04: 1.5428; 05: 1.72; 06: 1.62; 07: 1.55 | 22
| 1 | 6 | 01: 1.67; 02: 1.7273; 04: 1.4066; 05: 1.75; 06: 1.49; 07: 1.55
| Mean |  | 01: 1.615; 02: 1.7108; 04: 1.4659; 05: 1.7067; 06: 1.5567; 07: 1.5167
| Median |  | 01: 1.595; 02: 1.7147; 04: 1.4666; 05: 1.705; 06: 1.545; 07: 1.55
| Std Dev. |  | 01: 0.03937; 02: 0.024272; 04: 0.08063; 05: 0.03559; 06: 0.049666; 07: 0.08756
| Rel.Std.Dev. |  | 01: 2.438%; 02: 1.419%; 04: 5.5%; 05: 2.085%; 06: 3.191%; 07: 5.773%
| PDM3 |  | 01: 0.8848%; 02: 6.868%; 04: -8.429%; 05: 6.611%; 06: -2.759%; 07: -5.258%
| Z-Score (Absolute) |  | 01: 0.25; 02: 1.08; 04: 1.04; 05: 1.05; 06: 0.25; 07: 0.31
| Analytical results for Er in OREAS 604 (Certified Value 0.62 ppm) | [23 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 02; 02: 09; 04: 13; 05: 15; 06: 17; 07: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS | 2
3
| 1 | 1 | 01: 0.63; 02: 0.64; 04: 0.206; 05: 0.73; 06: 0.64; 07: 0.65 | 1
| 1 | 2 | 01: 0.61; 02: 0.61759; 04: 0.21517; 05: 0.64; 06: 0.62; 07: 0.6 | 0
| 1 | 3 | 01: 0.62; 02: 0.66122; 04: 0.20766; 05: 0.59; 06: 0.61; 07: 0.5 | 16
| 1 | 4 | 01: 0.62; 02: 0.61985; 04: 0.21975; 05: 0.67; 06: 0.61; 07: 0.55 | 0.62428
| 1 | 5 | 01: 0.66; 02: 0.6383; 04: 0.20015; 05: 0.6; 06: 0.66; 07: 0.55 | 23
| 1 | 6 | 01: 0.66; 02: 0.64144; 04: 0.18931; 05: 0.69; 06: 0.6; 07: 0.6
| Mean |  | 01: 0.63333; 02: 0.6364; 04: 0.20634; 05: 0.65333; 06: 0.62333; 07: 0.575
| Median |  | 01: 0.625; 02: 0.63915; 04: 0.20683; 05: 0.655; 06: 0.615; 07: 0.575
| Std Dev. |  | 01: 0.021602; 02: 0.016036; 04: 0.010844; 05: 0.053914; 06: 0.022509; 07: 0.05244
| Rel.Std.Dev. |  | 01: 3.411%; 02: 2.52%; 04: 5.255%; 05: 8.252%; 06: 3.611%; 07: 9.12%
| PDM3 |  | 01: 1.45%; 02: 1.942%; 04: -66.95%; 05: 4.654%; 06: -0.1517%; 07: -7.894%
| Z-Score (Absolute) |  | 01: 0.2; 02: 0.33; 04: 17.21; 05: 1.02; 06: 0.2; 07: 2.18
| Analytical results for Eu in OREAS 604 (Certified Value 0.74 ppm) | [23 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 02; 02: 09; 04: 13; 05: 15; 06: 17; 07: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS | 2
2
| 1 | 1 | 01: 0.61; 02: 0.63575; 04: -; 05: 0.86; 06: 0.84; 07: 0.7 | 1
| 1 | 2 | 01: 0.6; 02: 0.62623; 04: -; 05: 0.8; 06: 0.87; 07: 0.7 | 0
| 1 | 3 | 01: 0.6; 02: 0.65826; 04: -; 05: 0.82; 06: 0.85; 07: 0.75 | 16
| 1 | 4 | 01: 0.61; 02: 0.65327; 04: -; 05: 0.81; 06: 1.04; 07: 0.75 | 0.73786
| 1 | 5 | 01: 0.64; 02: 0.66238; 04: -; 05: 0.76; 06: 1.41; 07: 0.75 | 24
| 1 | 6 | 01: 0.61; 02: 0.64183; 04: -; 05: 0.85; 06: 0.89; 07: 0.65
| Mean |  | 01: 0.61167; 02: 0.64629; 05: 0.81667; 06: 0.98333; 07: 0.71667
| Median |  | 01: 0.61; 02: 0.64755; 05: 0.815; 06: 0.88; 07: 0.725
| Std Dev. |  | 01: 0.01472; 02: 0.014029; 05: 0.036148; 06: 0.22142; 07: 0.040825
| Rel.Std.Dev. |  | 01: 2.406%; 02: 2.171%; 05: 4.426%; 06: 22.52%; 07: 5.696%
| PDM3 |  | 01: -17.1%; 02: -12.41%; 05: 10.68%; 06: 33.27%; 07: -2.872%
| Z-Score (Absolute) |  | 01: 0.71; 02: 0.47; 04: NA; 05: 0.67; 06: 1.22; 07: 0
| Analytical results for Fe in OREAS 604 (Certified Value 3.02 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES |  | 2
3
| 1 | 1 | 2.9043 | 3 | 3.1 | 2.87 | 3 | 3 | 3.2 | 3.23 | 3.31 | 2.84 | 3.1464 | 3.03 | 3.16 | 2.65 | 2.97 | 3.09 | 2.785 | 2.79 | 2.73 | 3.25 | 3.24 |  | 1
| 1 | 2 | 2.9862 | 3 | 3.08 | 2.97 | 2.98 | 3.01 | 3.1707 | 3.27 | 3.34 | 2.69 | 3.1901 | 3.05 | 3.15 | 2.63 | 2.91 | 3.34 | 2.749 | 2.78 | 2.85 | 3.2 | 3.27 |  | 0
| 1 | 3 | 3.003 | 2.99 | 3.04 | 2.93 | 3.01 | 2.8 | 3.1706 | 3.28 | 3.32 | 2.76 | 3.1969 | 3.06 | 3.16 | 2.52 | 3.14 | 3.32 | 2.842 | 2.82 | 2.81 | 3.23 | 3.24 |  | 16
| 1 | 4 | 2.9883 | 3.01 | 3.06 | 2.91 | 2.96 | 2.96 | 3.1572 | 3.25 | 3.35 | 2.66 | 3.2005 | 3.06 | 3.15 | 2.49 | 3.01 | 3.32 | 2.867 | 2.69 | 2.76 | 3.32 | 3.22 |  | 3.0216
| 1 | 5 | 2.9799 | 3.07 | 3.01 | 3.09 | 2.91 | 2.94 | 3.1728 | 3.17 | 3.31 | 2.58 | 3.1243 | 3.05 | 3.12 | 2.5 | 3.05 | 3.28 | 2.739 | 2.83 | 2.76 | 3.32 | 3.21 |  | 25
| 1 | 6 | 3.0303 | 3.02 | 3.02 | 2.93 | 3.04 | 2.97 | 3.1471 | 3.25 | 3.32 | 2.68 | 3.1609 | 3.04 | 3.14 | 2.58 | 3.1 | 3.19 | 2.842 | 2.86 | 2.68 | 3.35 | 3.2
| Mean |  | 2.982 | 3.015 | 3.0517 | 2.95 | 2.9833 | 2.9467 | 3.1697 | 3.2417 | 3.325 | 2.7017 | 3.1698 | 3.0483 | 3.1467 | 2.5617 | 3.03 | 3.2567 | 2.804 | 2.795 | 2.765 | 3.2783 | 3.23
| Median |  | 2.9872 | 3.005 | 3.05 | 2.93 | 2.99 | 2.965 | 3.1706 | 3.25 | 3.32 | 2.685 | 3.1755 | 3.05 | 3.15 | 2.55 | 3.03 | 3.3 | 2.8135 | 2.805 | 2.76 | 3.285 | 3.23
| Std Dev. |  | 0.042126 | 0.02881 | 0.034881 | 0.075895 | 0.045019 | 0.076333 | 0.017872 | 0.0392 | 0.016432 | 0.089088 | 0.030935 | 0.01169 | 0.015055 | 0.068532 | 0.084617 | 0.097707 | 0.053792 | 0.058907 | 0.059582 | 0.059805 | 0.025298
| Rel.Std.Dev. |  | 1.413% | 0.9555% | 1.143% | 2.573% | 1.509% | 2.59% | 0.5638% | 1.209% | 0.4942% | 3.298% | 0.9759% | 0.3835% | 0.4785% | 2.675% | 2.793% | 3% | 1.918% | 2.108% | 2.155% | 1.824% | 0.7832%
| PDM3 |  | -1.311% | -0.2184% | 0.9951% | -2.37% | -1.266% | -2.48% | 4.902% | 7.283% | 10.04% | -10.59% | 4.906% | 0.8848% | 4.139% | -15.22% | 0.2781% | 7.78% | -7.201% | -7.499% | -8.492% | 8.497% | 6.897%
| Z-Score (Absolute) |  | 0.23 | 0.07 | 0.1 | 0.52 | 0.23 | 0.26 | 0.67 | 1.02 | 1.42 | 1.58 | 0.67 | 0.09 | 0.56 | 2.26 | 0 | 1.09 | 1.09 | 1.13 | 1.28 | 1.2 | 0.96
| Analytical results for Ga in OREAS 604 (Certified Value 26.7 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [6 empty] | 1
2
| 1 | 1 | 26.91 | 24.94 | 30 | 20 | 30 | 30 | 26.7791 | 29 | 17.92 | 22.6748 | 26.4 | 27.44 | 25.1 | 23.941 | 25.6 | 25.6 | [6 empty] | 1
| 1 | 2 | 26.864 | 24.55 | 30 | 30 | 30 | 30 | 26.9802 | 30 | 16.46 | 22.7897 | 26.7 | 25.29 | 24 | 23.585 | 25 | 25 | [6 empty] | 0
| 1 | 3 | 27.2525 | 24.67 | 30 | 30 | 30 | 20 | 26.7765 | 30 | 15.5 | 23.0304 | 26.4 | 27.05 | 25.2 | 23.777 | 25 | 25 | [6 empty] | 16
| 1 | 4 | 26.9875 | 25.12 | 30 | 30 | 30 | 20 | 27.1892 | 31 | 16.91 | 23.1654 | 26.8 | 25.12 | 24.7 | 24.539 | 26.8 | 25.4 | [6 empty] | 26.7033
| 1 | 5 | 26.838 | 25.68 | 30 | 30 | 20 | 30 | 26.947 | 31 | 15.28 | 22.1813 | 26.3 | 26.78 | 25.9 | 23.592 | 24.8 | 25.4 | [6 empty] | 26
| 1 | 6 | 26.5225 | 24.74 | 30 | 30 | 30 | 30 | 27.0082 | 30 | 16.41 | 23.0407 | 26.5 | 25.82 | 23.7 | 23.819 | 26.2 | 25.2
| Mean |  | 26.8957 | 24.95 | 30 | 28.3333 | 28.3333 | 26.6667 | 26.9467 | 30.1667 | 16.4133 | 22.8137 | 26.5167 | 26.25 | 24.7667 | 23.8755 | 25.5667 | 25.2667
| Median |  | 26.887 | 24.84 | 30 | 30 | 30 | 30 | 26.9636 | 30 | 16.435 | 22.91 | 26.45 | 26.3 | 24.9 | 23.798 | 25.3 | 25.3
| Std Dev. |  | 0.23645 | 0.41095 | 0 | 4.0825 | 4.0825 | 5.164 | 0.15544 | 0.75277 | 0.96307 | 0.35833 | 0.19408 | 0.97164 | 0.81404 | 0.35275 | 0.79415 | 0.24221
| Rel.Std.Dev. |  | 0.8791% | 1.647% | 0% | 14.41% | 14.41% | 19.36% | 0.5768% | 2.495% | 5.868% | 1.571% | 0.7319% | 3.701% | 3.287% | 1.477% | 3.106% | 0.9586%
| PDM3 |  | 0.7208% | -6.566% | 12.35% | 6.104% | 6.104% | -0.1371% | 0.9117% | 12.97% | -38.53% | -14.57% | -0.6988% | -1.697% | -7.252% | -10.59% | -4.256% | -5.38%
| Z-Score (Absolute) |  | 0.23 | 0.63 | 1.6 | 1.6 | 1.6 | 0.13 | 0.25 | 1.67 | 4.41 | 1.58 | 0.06 | 0.06 | 0.71 | 1.17 | 0.36 | 0.49
| Analytical results for Gd in OREAS 604 (Certified Value 2.72 ppm) | [23 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  | 2
2
| 1 | 1 | 01: 2.42; 02: 3.2262; 05: 2.4963; 06: 2.58; 07: 2.66; 08: 2.6 |  | 1
| 1 | 2 | 01: 2.5; 02: 3.2098; 05: 2.5199; 06: 2.73; 07: 2.61; 08: 2.2 |  | 0
| 1 | 3 | 01: 2.51; 02: 3.2353; 05: 2.5641; 06: 2.59; 07: 2.63; 08: 2.2 |  | 16
| 1 | 4 | 01: 2.58; 02: 3.2547; 05: 2.5821; 06: 2.59; 07: 2.55; 08: 2.2 |  | 2.7195
| 1 | 5 | 01: 2.65; 02: 3.4122; 05: 2.5578; 06: 2.52; 07: 2.72; 08: 2.2 |  | 27
| 1 | 6 | 01: 2.57; 02: 3.3544; 05: 2.5122; 06: 2.69; 07: 2.56; 08: 2.2
| Mean |  | 01: 2.5383; 02: 3.2821; 05: 2.5387; 06: 2.6167; 07: 2.6217; 08: 2.2667
| Median |  | 01: 2.54; 02: 3.245; 05: 2.5389; 06: 2.59; 07: 2.62; 08: 2.2
| Std Dev. |  | 01: 0.079352; 02: 0.081816; 05: 0.033899; 06: 0.077889; 07: 0.063692; 08: 0.1633
| Rel.Std.Dev. |  | 01: 3.126%; 02: 2.493%; 05: 1.335%; 06: 2.977%; 07: 2.429%; 08: 7.204%
| PDM3 |  | 01: -6.662%; 02: 20.69%; 05: -6.647%; 06: -3.781%; 07: -3.598%; 08: -16.65%
| Z-Score (Absolute) |  | 01: 0.64; 02: 11.4; 05: 0.63; 06: 0.63; 07: 0.71; 08: 6.11
| Analytical results for Ge in OREAS 604 (Certified Value < 5 ppm) | [23 empty] | Certified
Ge | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 14; 08: 15; 09: 17; 10: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: PF*MS |  | 2
2
| 1 | 1 | 01: 2.35; 02: 2.946; 05: < 10; 06: 1.78; 07: 3.2; 08: 1.43; 09: 0.5; 10: < 20 |  | 1
| 1 | 2 | 01: 2.45; 02: 3.2967; 05: < 10; 06: 1.85; 07: 3.6; 08: 1.31; 09: 0.2; 10: < 20 |  | 7
| 1 | 3 | 01: 2.26; 02: 2.9895; 05: < 10; 06: 1.72; 07: 3.2; 08: 1.52; 09: 0.2; 10: < 20 |  | 16
| 1 | 4 | 01: 2.19; 02: 3.0668; 05: < 10; 06: 1.93; 07: 3.3; 08: 1.35; 09: 0.2; 10: < 20 |  | < 5
| 1 | 5 | 01: 2.39; 02: 2.9616; 05: < 10; 06: 1.92; 07: 3.4; 08: 1.46; 09: < 0.1; 10: < 20 |  | 28
| 1 | 6 | 01: 2.32; 02: 3.3526; 05: < 10; 06: 1.82; 07: 3.2; 08: 1.45; 09: 0.3; 10: < 20
| Mean |  | 01: 2.3267; 02: 3.1022; 06: 1.8367; 07: 3.3167; 08: 1.42; 09: 0.28
| Median |  | 01: 2.335; 02: 3.0282; 06: 1.835; 07: 3.25; 08: 1.44; 09: 0.2
| Std Dev. |  | 01: 0.092664; 02: 0.17813; 06: 0.081158; 07: 0.16021; 08: 0.076942; 09: 0.13038
| Rel.Std.Dev. |  | 01: 3.983%; 02: 5.742%; 06: 4.419%; 07: 4.83%; 08: 5.418%; 09: 46.57%
PDM3
| Z-Score (Absolute) |  | 01: 0.2; 02: 0.82; 05: NA; 06: 0.2; 07: 0.94; 08: 0.53; 09: 1.52; 10: NA
| Analytical results for Hf in OREAS 604 (Certified Value 3.15 ppm) | [23 empty] | Certified
Hf | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 24; 12: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS |  | 2
3
| 1 | 1 | 01: 3.36; 02: 3.2542; 05: 2.9; 06: 3.155; 07: 3.5; 08: 3.29; 09: 3; 10: 3.306; 11: 3; 12: 2.8 |  | 1
| 1 | 2 | 01: 3.49; 02: 3.2254; 05: 2.6; 06: 3.2046; 07: 3.4; 08: 2.96; 09: 2.87; 10: 3.354; 11: 3; 12: 2.8 |  | 0
| 1 | 3 | 01: 3.38; 02: 3.3141; 05: 2.6; 06: 3.1631; 07: 3.4; 08: 3.18; 09: 3.09; 10: 3.113; 11: 3; 12: 3 |  | 16
| 1 | 4 | 01: 3.38; 02: 3.2761; 05: 2.9; 06: 3.2694; 07: 3.4; 08: 2.98; 09: 3.13; 10: 3.307; 11: 3.2; 12: 3 |  | 3.148
| 1 | 5 | 01: 3.35; 02: 3.2433; 05: 2.8; 06: 3.3139; 07: 3.2; 08: 3.4; 09: 3.57; 10: 3.221; 11: 3; 12: 3 |  | 29
| 1 | 6 | 01: 3.32; 02: 3.1707; 05: 2.8; 06: 3.2235; 07: 3.5; 08: 3.13; 09: 2.74; 10: 3.296; 11: 3.2; 12: 2.8
| Mean |  | 01: 3.38; 02: 3.2473; 05: 2.7667; 06: 3.2216; 07: 3.4; 08: 3.1567; 09: 3.0667; 10: 3.2662; 11: 3.0667; 12: 2.9
| Median |  | 01: 3.37; 02: 3.2487; 05: 2.8; 06: 3.214; 07: 3.4; 08: 3.155; 09: 3.045; 10: 3.301; 11: 3; 12: 2.9
| Std Dev. |  | 01: 0.05831; 02: 0.048359; 05: 0.13663; 06: 0.061593; 07: 0.10954; 08: 0.17212; 09: 0.28556; 10: 0.086439; 11: 0.10328; 12: 0.10954
| Rel.Std.Dev. |  | 01: 1.725%; 02: 1.489%; 05: 4.938%; 06: 1.912%; 07: 3.222%; 08: 5.453%; 09: 9.312%; 10: 2.647%; 11: 3.368%; 12: 3.777%
| PDM3 |  | 01: 7.369%; 02: 3.153%; 05: -12.11%; 06: 2.336%; 07: 8.004%; 08: 0.2743%; 09: -2.585%; 10: 3.753%; 11: -2.585%; 12: -7.879%
| Z-Score (Absolute) |  | 01: 0.93; 02: 0.32; 05: 2.33; 06: 0.18; 07: 1.16; 08: 0.18; 09: 0.67; 10: 0.59; 11: 0.67; 12: 1.59
| Analytical results for Hg in OREAS 604 (Indicative Value 0.67 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 11 |  | ppm
|  |  | 01: 4A*OES |  | 2
2
| 1 | 1 | [22 empty] | 1 | bar: 01
| 1 | 2 | 01: < 1 |  | 3
| 1 | 3 | 01: < 1 |  | 16
| 1 | 4 | 01: < 1 |  | 0.66667
| 1 | 5 | 01: < 1 |  | 9
| 1 | 6 | bar: 01
Mean | bar: 01
Median | bar: 01
| Std Dev. |  | 01: 0
Rel.Std.Dev. | bar: 01
| PDM3 |  | 01: 50%
Z-Score (Absolute) | bar: 01
| Analytical results for Ho in OREAS 604 (Certified Value 0.25 ppm) | [23 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  | 2
2
| 1 | 1 | 01: 0.27; 02: 0.27612; 05: 0.206; 06: 0.28; 07: 0.24; 08: 0.26 |  | 1
| 1 | 2 | 01: 0.26; 02: 0.26315; 05: 0.21517; 06: 0.27; 07: 0.23; 08: 0.2 |  | 0
| 1 | 3 | 01: 0.26; 02: 0.26976; 05: 0.20766; 06: 0.28; 07: 0.24; 08: 0.22 |  | 16
| 1 | 4 | 01: 0.27; 02: 0.26465; 05: 0.21975; 06: 0.28; 07: 0.23; 08: 0.22 |  | 0.2461
| 1 | 5 | 01: 0.28; 02: 0.26376; 05: 0.20015; 06: 0.26; 07: 0.26; 08: 0.2 |  | 30
| 1 | 6 | 01: 0.27; 02: 0.26426; 05: 0.19931; 06: 0.28; 07: 0.23; 08: 0.22
| Mean |  | 01: 0.26833; 02: 0.26695; 05: 0.20801; 06: 0.275; 07: 0.23833; 08: 0.22
| Median |  | 01: 0.27; 02: 0.26445; 05: 0.20683; 06: 0.28; 07: 0.235; 08: 0.22
| Std Dev. |  | 01: 0.0075277; 02: 0.0050815; 05: 0.0081328; 06: 0.0083666; 07: 0.01169; 08: 0.021909
| Rel.Std.Dev. |  | 01: 2.805%; 02: 1.904%; 05: 3.91%; 06: 3.042%; 07: 4.905%; 08: 9.959%
| PDM3 |  | 01: 9.033%; 02: 8.47%; 05: -15.48%; 06: 11.74%; 07: -3.157%; 08: -10.61%
| Z-Score (Absolute) |  | 01: 0.56; 02: 0.51; 05: 1.58; 06: 0.79; 07: 0.51; 08: 1.16
| Analytical results for In in OREAS 604 (Certified Value 3.57 ppm) | [23 empty] | Certified
In | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 12; 08: 13; 09: 14; 10: 15; 11: 17; 12: 19; 13: 24; 14: 25 |  | ppm
|  |  | 01: 4A*OES; 02: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*OES/MS; 12: 4A*OES/MS; 13: 4A*MS; 14: 4A*MS |  | 2
3
| 1 | 1 | 01: < 5; 02: 3.38; 05: 3.6842; 06: < 10; 07: 3.8; 08: 3.3552; 09: 3.9; 10: 4; 11: 3.92; 12: 3.07; 13: 3.8; 14: 3.32 |  | 1
| 1 | 2 | 01: < 5; 02: 3.35; 05: 3.8196; 06: < 10; 07: 3.58; 08: 3.3461; 09: 3.78; 10: 4.02; 11: 3.71; 12: 3.037; 13: 3.65; 14: 3.3 |  | 0
| 1 | 3 | 01: < 5; 02: 3.48; 05: 3.7216; 06: < 10; 07: 3.6; 08: 3.3264; 09: 3.79; 10: 3.83; 11: 3.86; 12: 3.022; 13: 3.65; 14: 3.3 |  | 16
| 1 | 4 | 01: < 5; 02: 3.4; 05: 3.6667; 06: < 10; 07: 3.7; 08: 3.3478; 09: 3.94; 10: 3.95; 11: 3.69; 12: 3.06; 13: 3.9; 14: 3.32 |  | 3.5749
| 1 | 5 | 01: < 5; 02: 3.46; 05: 3.6621; 06: < 10; 07: 3.27; 08: 3.2701; 09: 3.78; 10: 3.78; 11: 3.98; 12: 3.032; 13: 3.8; 14: 3.28 |  | 31
| 1 | 6 | 01: < 5; 02: 3.52; 05: 3.7056; 06: < 10; 07: 3.59; 08: 3.2025; 09: 3.68; 10: 4.03; 11: 3.89; 12: 3.039; 13: 3.85; 14: 3.3
| Mean |  | 02: 3.4317; 05: 3.71; 07: 3.59; 08: 3.308; 09: 3.8117; 10: 3.935; 11: 3.8417; 12: 3.0433; 13: 3.775; 14: 3.3033
| Median |  | 02: 3.43; 05: 3.6949; 07: 3.595; 08: 3.3362; 09: 3.785; 10: 3.975; 11: 3.875; 12: 3.038; 13: 3.8; 14: 3.3
| Std Dev. |  | 02: 0.065243; 05: 0.058308; 07: 0.17821; 08: 0.06026; 09: 0.093897; 10: 0.10559; 11: 0.11686; 12: 0.018063; 13: 0.10368; 14: 0.015055
| Rel.Std.Dev. |  | 02: 1.901%; 05: 1.572%; 07: 4.964%; 08: 1.822%; 09: 2.463%; 10: 2.683%; 11: 3.042%; 12: 0.5935%; 13: 2.747%; 14: 0.4558%
| PDM3 |  | 02: -4.006%; 05: 3.779%; 07: 0.4228%; 08: -7.465%; 09: 6.623%; 10: 10.07%; 11: 7.463%; 12: -14.87%; 13: 5.598%; 14: -7.596%
| Z-Score (Absolute) |  | 01: NA; 02: 0.68; 05: 0.16; 06: NA; 07: 0.16; 08: 1.02; 09: 0.57; 10: 0.97; 11: 0.67; 12: 1.96; 13: 0.45; 14: 1.1
| Analytical results for K in OREAS 604 (Certified Value 1.32 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES |  |  | 2
3
| 1 | 1 | 1.25 | 1.3453 | 1.31 | 1.29 | 1.3 | 1.25 | 1.2707 | 1.4 | 1.42 | 1.3523 | 1.3112 | 1.34 | 1.28 | 1.33 | 1.26 | 1.315 | 1.39 | 1.21 | 1.37 | 1.4 |  |  | 1
| 1 | 2 | 1.25 | 1.3595 | 1.32 | 1.32 | 1.29 | 1.25 | 1.2739 | 1.39 | 1.28 | 1.3535 | 1.3195 | 1.35 | 1.27 | 1.29 | 1.36 | 1.327 | 1.41 | 1.22 | 1.34 | 1.3 |  |  | 0
| 1 | 3 | 1.255 | 1.3545 | 1.28 | 1.31 | 1.29 | 1.2 | 1.2794 | 1.42 | 1.33 | 1.3605 | 1.3163 | 1.34 | 1.22 | 1.4 | 1.4 | 1.3 | 1.33 | 1.22 | 1.32 | 1.4 |  |  | 16
| 1 | 4 | 1.255 | 1.3512 | 1.29 | 1.32 | 1.27 | 1.23 | 1.2667 | 1.44 | 1.4 | 1.3761 | 1.3204 | 1.34 | 1.2 | 1.34 | 1.37 | 1.337 | 1.4 | 1.25 | 1.38 | 1.3 |  |  | 1.3181
| 1 | 5 | 1.22 | 1.3433 | 1.26 | 1.38 | 1.25 | 1.24 | 1.2613 | 1.43 | 1.38 | 1.3197 | 1.3244 | 1.33 | 1.21 | 1.37 | 1.33 | 1.304 | 1.33 | 1.22 | 1.4 | 1.4 |  |  | 32
| 1 | 6 | 1.24 | 1.3476 | 1.28 | 1.29 | 1.3 | 1.25 | 1.2721 | 1.36 | 1.42 | 1.358 | 1.3254 | 1.34 | 1.25 | 1.38 | 1.65 | 1.333 | 1.31 | 1.2 | 1.4 | 1.4
| Mean |  | 1.245 | 1.3502 | 1.29 | 1.3183 | 1.2833 | 1.2367 | 1.2707 | 1.4067 | 1.3717 | 1.3533 | 1.3195 | 1.34 | 1.2383 | 1.3517 | 1.395 | 1.3193 | 1.3617 | 1.22 | 1.3683 | 1.3667
| Median |  | 1.25 | 1.3494 | 1.285 | 1.315 | 1.29 | 1.245 | 1.2714 | 1.41 | 1.39 | 1.3558 | 1.3199 | 1.34 | 1.235 | 1.355 | 1.365 | 1.321 | 1.36 | 1.22 | 1.375 | 1.4
| Std Dev. |  | 0.013416 | 0.0060721 | 0.021909 | 0.033116 | 0.019664 | 0.019664 | 0.0062008 | 0.029439 | 0.056006 | 0.018576 | 0.0052679 | 0.0063246 | 0.033116 | 0.039707 | 0.13368 | 0.015397 | 0.043089 | 0.016733 | 0.032506 | 0.05164
| Rel.Std.Dev. |  | 1.078% | 0.4497% | 1.698% | 2.512% | 1.532% | 1.59% | 0.488% | 2.093% | 4.083% | 1.373% | 0.3992% | 0.472% | 2.674% | 2.938% | 9.583% | 1.167% | 3.164% | 1.372% | 2.376% | 3.779%
| PDM3 |  | -5.549% | 2.435% | -2.135% | 0.01468% | -2.641% | -6.181% | -3.601% | 6.716% | 4.061% | 2.671% | 0.1057% | 1.658% | -6.054% | 2.543% | 5.831% | 0.09055% | 3.302% | -7.445% | 3.808% | 3.681%
| Z-Score (Absolute) |  | 1.51 | 0.37 | 0.71 | 0.2 | 0.83 | 1.53 | 1.06 | 1.37 | 0.75 | 0.42 | 0.18 | 0.18 | 1.63 | 0.39 | 0.25 | 0.19 | 0.57 | 1.96 | 0.69 | 0.66
| Analytical results for La in OREAS 604 (Certified Value 19.4 ppm) | [23 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 23 | 25 | [9 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | [9 empty] | 1
1
| 1 | 1 | 21.49 | 21.56 | 20.0294 | 17 | 7.4 | 21.0444 | 23 | 10.2 | 17.6 | < 20 | 13.685 | 8 | 19.4 | [9 empty] | 1
| 1 | 2 | 22.48 | 20.59 | 20.2374 | 16 | 8 | 21.5119 | 23 | 8.4 | 17.4 | < 20 | 13.78 | 11 | 18.6 | [9 empty] | 0
| 1 | 3 | 22.685 | 20.96 | 20.514 | 15 | 6.3 | 20.6236 | 23 | 10 | 17 | < 20 | 13.66 | 10 | 20.6 | [9 empty] | 16
| 1 | 4 | 21.6 | 20.93 | 20.181 | 16 | 8.1 | 20.7864 | 23 | 10.7 | 19.5 | < 20 | 13.21 | 10 | 19.5 | [9 empty] | 19.3687
| 1 | 5 | 22.535 | 20.79 | 20.7834 | 15 | 4.9 | 21.5355 | 23 | 9.6 | 21.8 | < 20 | 12.81 | 11 | 21 | [9 empty] | 33
| 1 | 6 | 22.08 | 21.73 | 20.2593 | 16 | 8 | 20.7892 | 24 | 8.8 | 17.1 | < 20 | 13.42 | 7 | 18.2
| Mean |  | 22.145 | 21.0933 | 20.3341 | 15.8333 | 7.1167 | 21.0485 | 23.1667 | 9.6167 | 18.4 |  | 13.4275 | 9.5 | 19.55
| Median |  | 22.28 | 20.945 | 20.2484 | 16 | 7.7 | 20.9168 | 23 | 9.8 | 17.5 |  | 13.54 | 10 | 19.45
| Std Dev. |  | 0.50719 | 0.4501 | 0.27036 | 0.75277 | 1.2797 | 0.39203 | 0.40825 | 0.87274 | 1.9005 |  | 0.36696 | 1.6432 | 1.0913
| Rel.Std.Dev. |  | 2.29% | 2.134% | 1.33% | 4.754% | 17.98% | 1.862% | 1.762% | 9.075% | 10.33% |  | 2.733% | 17.3% | 5.582%
| PDM3 |  | 14.33% | 8.904% | 4.984% | -18.25% | -63.26% | 8.673% | 19.61% | -50.35% | -5.001% |  | -30.67% | -50.95% | 0.936%
| Z-Score (Absolute) |  | 0.75 | 0.53 | 0.36 | 0.6 | 2.37 | 0.52 | 0.97 | 1.93 | 0.2 | NA | 1.11 | 1.95 | 0.2
| Analytical results for Li in OREAS 604 (Certified Value 22.6 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | [7 empty] | 1
2
| 1 | 1 | 20.46 | 22.1 | 26.4595 | 24 | 25 | 25.3086 | 21.8 | 23 | 22 | 22 | 33 | 20.133 | 47 | 20 | 22.5 | [7 empty] | 1
| 1 | 2 | 21.49 | 22.5 | 25.6529 | 24 | 23.7 | 25.2349 | 21.4 | 24 | 22 | 22 | 36 | 20.089 | 48 | 20 | 23 | [7 empty] | 0
| 1 | 3 | 21.0575 | 22 | 25.0451 | 23 | 24.2 | 25.689 | 21.7 | 25 | 21 | 23 | 37 | 19.874 | 49 | 20 | 23 | [7 empty] | 16
| 1 | 4 | 21.13 | 22.9 | 26.1638 | 24 | 24.2 | 25.6684 | 21.4 | 23 | 21 | 23 | 36 | 20.398 | 48 | 20 | 22.5 | [7 empty] | 22.6336
| 1 | 5 | 20.7425 | 22.2 | 25.5713 | 24 | 22.9 | 25.2809 | 21.7 | 23 | 21 | 23 | 37 | 19.561 | 51 | 30 | 23 | [7 empty] | 34
| 1 | 6 | 21 | 22 | 26.9484 | 23 | 24.7 | 25.2356 | 21.4 | 23 | 21 | 23 | 35 | 20.167 | 51 | 30 | 23
| Mean |  | 20.98 | 22.2833 | 25.9735 | 23.6667 | 24.1167 | 25.4029 | 21.5667 | 23.5 | 21.3333 | 22.6667 | 35.6667 | 20.037 | 49 | 23.3333 | 22.8333
| Median |  | 21.0287 | 22.15 | 25.9084 | 24 | 24.2 | 25.2947 | 21.55 | 23 | 21 | 23 | 36 | 20.111 | 48.5 | 20 | 23
| Std Dev. |  | 0.35095 | 0.35449 | 0.68548 | 0.5164 | 0.74677 | 0.21556 | 0.18619 | 0.83666 | 0.5164 | 0.5164 | 1.5055 | 0.28715 | 1.6733 | 5.164 | 0.2582
| Rel.Std.Dev. |  | 1.673% | 1.591% | 2.639% | 2.182% | 3.096% | 0.8486% | 0.8633% | 3.56% | 2.421% | 2.278% | 4.221% | 1.433% | 3.415% | 22.13% | 1.131%
| PDM3 |  | -7.306% | -1.548% | 14.76% | 4.564% | 6.553% | 12.24% | -4.714% | 3.828% | -5.745% | 0.1461% | 57.58% | -11.47% | 116.5% | 3.092% | 0.8825%
| Z-Score (Absolute) |  | 0.83 | 0.3 | 1.41 | 0.37 | 0.58 | 1.16 | 0.57 | 0.3 | 0.67 | 0.07 | 5.77 | 1.26 | 11.76 | 1.27 | 0
| Analytical results for Lu in OREAS 604 (Certified Value 81 ppb) | [23 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab | [7 empty] | 1
| No | No | 01: 02; 02: 09; 09: 15; 11: 17; 12: 19; 13: 25 | [7 empty] | ppb
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 09: 4A*MS; 11: 4A*OES/MS; 12: 4A*OES/MS; 13: 4A*MS | [7 empty] | 0
0
| 1 | 1 | 01: 70; 02: 85.4132; 09: 100; 11: 70; 12: 87; 13: 100 | [7 empty] | 1
| 1 | 2 | 01: 80; 02: 83.8521; 09: 100; 11: 70; 12: 83; 13: 80 | [7 empty] | 0
| 1 | 3 | 01: 70; 02: 85.513; 09: < 100; 11: 70; 12: 85; 13: 80 | [7 empty] | 16
| 1 | 4 | 01: 80; 02: 90.8943; 09: 100; 11: 70; 12: 89; 13: 80 | [7 empty] | 80.7266
| 1 | 5 | 01: 70; 02: 89.2321; 09: 100; 11: 80; 12: 84; 13: 100 | [7 empty] | 35
| 1 | 6 | 01: 70; 02: 84.2511; 09: < 100; 11: 70; 12: 89; 13: 60
| Mean |  | 01: 73.3333; 02: 86.526; 09: 100; 11: 71.6667; 12: 86.1667; 13: 83.3333
| Median |  | 01: 70; 02: 85.4631; 09: 100; 11: 70; 12: 86; 13: 80
| Std Dev. |  | 01: 5.164; 02: 2.8634; 09: 0; 11: 4.0825; 12: 2.5626; 13: 15.0555
| Rel.Std.Dev. |  | 01: 7.042%; 02: 3.309%; 09: 0%; 11: 5.696%; 12: 2.974%; 13: 18.07%
| PDM3 |  | 01: -9.158%; 02: 7.184%; 09: 23.87%; 11: -11.22%; 12: 6.739%; 13: 3.229%
| Z-Score (Absolute) |  | 01: 2.24; 02: 0.71; 09: 0; 11: 2.61; 12: 0.63; 13: 0
| Analytical results for Mg in OREAS 604 (Certified Value 0.208 wt.%) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | PF*OES |  |  | 3
3
| 1 | 1 | 0.19144 | 0.2024 | 0.2 | 0.2 | 0.2 | 0.19 | 0.20648 | 0.22 | 0.22 | 0.23 | 0.21063 | 0.204 | 0.23 | 0.2 | 0.21 | 0.19 | 0.196 | 0.19 | 0.21 | 0.24 |  |  | 1
| 1 | 2 | 0.19728 | 0.1984 | 0.2 | 0.2 | 0.2 | 0.19 | 0.20566 | 0.225 | 0.22 | 0.21 | 0.21139 | 0.2028 | 0.23 | 0.2 | 0.21 | 0.21 | 0.195 | 0.2 | 0.21 | 0.25 |  |  | 0
| 1 | 3 | 0.19528 | 0.2015 | 0.19 | 0.2 | 0.2 | 0.18 | 0.20325 | 0.22 | 0.22 | 0.22 | 0.21288 | 0.2057 | 0.23 | 0.19 | 0.22 | 0.21 | 0.199 | 0.2 | 0.21 | 0.24 |  |  | 16
| 1 | 4 | 0.19552 | 0.2015 | 0.2 | 0.2 | 0.19 | 0.19 | 0.20481 | 0.22 | 0.22 | 0.22 | 0.21668 | 0.2072 | 0.23 | 0.19 | 0.22 | 0.21 | 0.21 | 0.2 | 0.21 | 0.23 |  |  | 0.20768
| 1 | 5 | 0.19304 | 0.2057 | 0.19 | 0.21 | 0.19 | 0.19 | 0.20611 | 0.22 | 0.23 | 0.21 | 0.20635 | 0.2027 | 0.23 | 0.19 | 0.22 | 0.21 | 0.193 | 0.2 | 0.22 | 0.25 |  |  | 36
| 1 | 6 | 0.19728 | 0.2029 | 0.19 | 0.2 | 0.2 | 0.19 | 0.20922 | 0.22 | 0.22 | 0.23 | 0.21265 | 0.2054 | 0.23 | 0.2 | 0.21 | 0.21 | 0.198 | 0.19 | 0.22 | 0.23
| Mean |  | 0.19497 | 0.20207 | 0.195 | 0.20167 | 0.19667 | 0.18833 | 0.20592 | 0.22083 | 0.22167 | 0.22 | 0.21176 | 0.20463 | 0.23 | 0.195 | 0.215 | 0.20667 | 0.1985 | 0.19667 | 0.21333 | 0.24
| Median |  | 0.1954 | 0.20195 | 0.195 | 0.2 | 0.2 | 0.19 | 0.20588 | 0.22 | 0.22 | 0.22 | 0.21202 | 0.2047 | 0.23 | 0.195 | 0.215 | 0.21 | 0.197 | 0.2 | 0.21 | 0.24
| Std Dev. |  | 0.0023348 | 0.0023704 | 0.0054772 | 0.0040825 | 0.005164 | 0.0040825 | 0.0019817 | 0.0020412 | 0.0040825 | 0.0089443 | 0.0033752 | 0.0017784 | 0 | 0.0054772 | 0.0054772 | 0.008165 | 0.0060249 | 0.005164 | 0.005164 | 0.0089443
| Rel.Std.Dev. |  | 1.197% | 1.173% | 2.809% | 2.024% | 2.626% | 2.168% | 0.9624% | 0.9243% | 1.842% | 4.066% | 1.594% | 0.8691% | 0% | 2.809% | 2.548% | 3.951% | 3.035% | 2.626% | 2.421% | 3.727%
| PDM3 |  | -6.12% | -2.705% | -6.108% | -2.898% | -5.305% | -9.318% | -0.8492% | 6.331% | 6.732% | 5.93% | 1.964% | -1.469% | 10.74% | -6.108% | 3.522% | -0.4901% | -4.422% | -5.305% | 2.72% | 15.56%
| Z-Score (Absolute) |  | 0.76 | 0.24 | 0.76 | 0.27 | 0.63 | 1.25 | 0.05 | 1.14 | 1.21 | 1.08 | 0.48 | 0.05 | 1.82 | 0.76 | 0.72 | 0.1 | 0.5 | 0.63 | 0.59 | 2.19
| Analytical results for Mn in OREAS 604 (Certified Value 244 ppm) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 234.955 | 238 | 250 | 247 | 249 | 245 | 241.991 | 245 | 260 | 205 | 249.234 | 253 | 253 | 222 | 233 | 200 | 250.519 | 263 | 214 | 228 | 240 |  | 1
| 1 | 2 | 238.52 | 238 | 251 | 256 | 243 | 233 | 244.6201 | 248 | 267 | 188 | 255.09 | 251 | 257 | 221 | 235 | 300 | 246.793 | 260 | 236 | 226 | 238 |  | 0
| 1 | 3 | 239.4567 | 237 | 245 | 250 | 249 | 219 | 242.5102 | 250 | 269 | 198 | 252.999 | 254 | 251 | 215 | 236 | 300 | 240.033 | 264 | 229 | 234 | 242 |  | 16
| 1 | 4 | 239.65 | 240 | 245 | 253 | 245 | 233 | 242.7497 | 249 | 273 | 196 | 259.107 | 253 | 250 | 208 | 235 | 300 | 262.967 | 253 | 237 | 236 | 236 |  | 243.748
| 1 | 5 | 239.34 | 226 | 240 | 266 | 238 | 240 | 243.2309 | 243 | 271 | 201 | 245.413 | 248 | 254 | 216 | 242 | 300 | 232.797 | 264 | 235 | 244 | 238 |  | 37
| 1 | 6 | 240.3025 | 235 | 243 | 252 | 249 | 233 | 243.2181 | 247 | 262 | 203 | 254.504 | 252 | 258 | 218 | 229 | 300 | 234.475 | 268 | 213 | 252 | 226
| Mean |  | 238.704 | 235.6667 | 245.6667 | 254 | 245.5 | 233.8333 | 243.0533 | 247 | 267 | 198.5 | 252.7245 | 251.8333 | 253.8333 | 216.6667 | 235 | 283.3333 | 244.5973 | 262 | 227.3333 | 236.6667 | 236.6667
| Median |  | 239.3983 | 237.5 | 245 | 252.5 | 247 | 233 | 242.9839 | 247.5 | 268 | 199.5 | 253.7515 | 252.5 | 253.5 | 217 | 235 | 300 | 243.413 | 263.5 | 232 | 235 | 238
| Std Dev. |  | 1.9241 | 5.0067 | 4.1793 | 6.603 | 4.4609 | 8.7731 | 0.89768 | 2.6077 | 5.099 | 6.091 | 4.8003 | 2.137 | 3.1885 | 5.0465 | 4.2426 | 40.8248 | 11.3126 | 5.099 | 11.0755 | 9.8522 | 5.6095
| Rel.Std.Dev. |  | 0.806% | 2.124% | 1.701% | 2.6% | 1.817% | 3.752% | 0.3693% | 1.056% | 1.91% | 3.069% | 1.899% | 0.8486% | 1.256% | 2.329% | 1.805% | 14.41% | 4.625% | 1.946% | 4.872% | 4.163% | 2.37%
| PDM3 |  | -2.069% | -3.315% | 0.7871% | 4.206% | 0.7188% | -4.068% | -0.285% | 1.334% | 9.539% | -18.56% | 3.683% | 3.317% | 4.138% | -11.11% | -3.589% | 16.24% | 0.3484% | 7.488% | -6.734% | -2.905% | -2.905%
| Z-Score (Absolute) |  | 0.5 | 0.59 | 0.09 | 0.6 | 0.08 | 0.92 | 0.13 | 0.2 | 1.9 | 3.92 | 0.69 | 0.62 | 0.79 | 2.37 | 0.82 | 3.29 | 0 | 1.63 | 1.47 | 0.67 | 0.49
| Analytical results for Mo in OREAS 604 (Certified Value 4.11 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 4.5046 | 3.9 | 4 | 4 | 3 | 4 | 4.3063 | 5 | 5 | 3.88 | 4.3 | 5 | 11 | 5.73 | 4 | 4.094 | 4 | 4 | 4.5 |  |  |  | 1
| 1 | 2 | 4.2161 | 3.8 | 3 | 4 | 3 | 3 | 4.0498 | 4 | 5 | 2.81 | 4.3 | 4 | 12 | 5.13 | 4 | 3.982 | 4 | 4 | 4.5 |  |  |  | 0
| 1 | 3 | 4.0584 | 3.8 | 4 | 4 | 3 | 4 | 4.1199 | 4 | 6 | 3.09 | 4.4 | 4 | 12 | 5.16 | 4 | 4.158 | 4 | 4.5 | 4.5 |  |  |  | 16
| 1 | 4 | 4.3833 | 3.7 | 4 | 4 | 3 | 3 | 4.092 | 4 | 5 | 2.94 | 4.3 | 4 | 12 | 4.87 | 5 | 4.154 | 5 | 4.5 | 4.5 |  |  |  | 4.1145
| 1 | 5 | 4.3097 | 3.8 | 4 | 4 | 3 | 4 | 3.943 | 4 | 5 | 3.28 | 4.3 | 5 | 10 | 5.25 | 3 | 4.091 | 4 | 4 | 4.5 |  |  |  | 38
| 1 | 6 | 4.3962 | 3.9 | 3 | 4 | 4 | 4 | 4.498 | 4 | 5 | 3.59 | 4.4 | 4 | 10 | 4.71 | 4 | 4.175 | 4 | 4.5 | 4.5
| Mean |  | 4.3114 | 3.8167 | 3.6667 | 4 | 3.1667 | 3.6667 | 4.1682 | 4.1667 | 5.1667 | 3.265 | 4.3333 | 4.3333 | 11.1667 | 5.1417 | 4 | 4.109 | 4.1667 | 4.25 | 4.5
| Median |  | 4.3465 | 3.8 | 4 | 4 | 3 | 4 | 4.106 | 4 | 5 | 3.185 | 4.3 | 4 | 11.5 | 5.145 | 4 | 4.124 | 4 | 4.25 | 4.5
| Std Dev. |  | 0.15669 | 0.075277 | 0.5164 | 0 | 0.40825 | 0.5164 | 0.20035 | 0.40825 | 0.40825 | 0.40678 | 0.05164 | 0.5164 | 0.98319 | 0.35171 | 0.63246 | 0.071358 | 0.40825 | 0.27386 | 0
| Rel.Std.Dev. |  | 3.634% | 1.972% | 14.08% | 0% | 12.89% | 14.08% | 4.807% | 9.798% | 7.902% | 12.46% | 1.192% | 11.92% | 8.805% | 6.84% | 15.81% | 1.737% | 9.798% | 6.444% | 0%
| PDM3 |  | 4.784% | -7.24% | -10.89% | -2.784% | -23.04% | -10.89% | 1.303% | 1.267% | 25.57% | -20.65% | 5.317% | 5.317% | 171.4% | 24.96% | -2.784% | -0.1348% | 1.267% | 3.292% | 9.368%
| Z-Score (Absolute) |  | 0.59 | 1.42 | 2.02 | 0.67 | 4.05 | 2.02 | 0.01 | 0 | 3.37 | 3.65 | 0.67 | 0.67 | 28.32 | 3.94 | 0.67 | 0.23 | 0 | 0.34 | 1.35
| Analytical results for Na in OREAS 604 (Certified Value 0.836 wt.%) | [23 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.7741 | 0.8349 | 0.82 | 0.82 | 0.82 | 0.81 | 0.89569 | 0.84 | 0.89 | 0.81997 | 0.8138 | 0.79 | 0.8 | 0.84 | 0.8 | 0.842 | 0.73 | 0.88 | 0.87 |  |  |  | 1
| 1 | 2 | 0.8002 | 0.8594 | 0.84 | 0.84 | 0.81 | 0.79 | 0.88478 | 0.83 | 0.84 | 0.83166 | 0.8177 | 0.81 | 0.8 | 0.84 | 0.87 | 0.834 | 0.74 | 0.87 | 0.87 |  |  |  | 0
| 1 | 3 | 0.7969 | 0.8439 | 0.81 | 0.84 | 0.82 | 0.78 | 0.8887 | 0.84 | 0.89 | 0.82497 | 0.8195 | 0.81 | 0.77 | 0.86 | 0.87 | 0.823 | 0.74 | 0.86 | 0.88 |  |  |  | 16
| 1 | 4 | 0.79165 | 0.8511 | 0.82 | 0.84 | 0.81 | 0.83 | 0.8848 | 0.83 | 0.85 | 0.83377 | 0.8173 | 0.8 | 0.76 | 0.83 | 0.87 | 0.839 | 0.76 | 0.89 | 0.89 |  |  |  | 0.83567
| 1 | 5 | 0.77915 | 0.8304 | 0.81 | 0.87 | 0.79 | 0.82 | 0.89292 | 0.83 | 0.86 | 0.82497 | 0.8205 | 0.8 | 0.76 | 0.84 | 0.88 | 0.83 | 0.74 | 0.9 | 0.88 |  |  |  | 39
| 1 | 6 | 0.7919 | 0.8311 | 0.81 | 0.83 | 0.83 | 0.83 | 0.89936 | 0.83 | 0.89 | 0.83042 | 0.821 | 0.81 | 0.79 | 0.85 | 0.84 | 0.842 | 0.73 | 0.91 | 0.87
| Mean |  | 0.78898 | 0.8418 | 0.81833 | 0.84 | 0.81333 | 0.81 | 0.89104 | 0.83333 | 0.87 | 0.82763 | 0.8183 | 0.80333 | 0.78 | 0.84333 | 0.855 | 0.835 | 0.74 | 0.885 | 0.87667
| Median |  | 0.79178 | 0.8394 | 0.815 | 0.84 | 0.815 | 0.815 | 0.89081 | 0.83 | 0.875 | 0.8277 | 0.8186 | 0.805 | 0.78 | 0.84 | 0.87 | 0.8365 | 0.74 | 0.885 | 0.875
| Std Dev. |  | 0.010221 | 0.01177 | 0.01169 | 0.016733 | 0.013663 | 0.020976 | 0.0059688 | 0.005164 | 0.022804 | 0.0051877 | 0.0026525 | 0.008165 | 0.018974 | 0.010328 | 0.030166 | 0.0075366 | 0.010954 | 0.018708 | 0.008165
| Rel.Std.Dev. |  | 1.295% | 1.398% | 1.429% | 1.992% | 1.68% | 2.59% | 0.6699% | 0.6197% | 2.621% | 0.6268% | 0.3242% | 1.016% | 2.433% | 1.225% | 3.528% | 0.9026% | 1.48% | 2.114% | 0.9314%
| PDM3 |  | -5.587% | 0.7334% | -2.075% | 0.518% | -2.673% | -3.072% | 6.626% | -0.2798% | 4.108% | -0.9626% | -2.079% | -3.87% | -6.662% | 0.9168% | 2.313% | -0.08035% | -11.45% | 5.903% | 4.906%
| Z-Score (Absolute) |  | 1.28 | 0.24 | 0.43 | 0.19 | 0.58 | 0.67 | 1.67 | 0 | 1.06 | 0.16 | 0.43 | 0.87 | 1.54 | 0.29 | 0.94 | 0.05 | 2.7 | 1.49 | 1.25
| Analytical results for Nb in OREAS 604 (Certified Value 8.82 ppm) | [23 empty] | Certified
Nb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 05: 09; 06: 11; 07: 12; 08: 14; 09: 15; 11: 17; 12: 19; 13: 23; 14: 25 |  |  |  | ppm
|  |  | 01: 4A*OES; 02: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 11: 4A*OES/MS; 12: 4A*OES/MS; 13: 4A*OES; 14: 4A*MS |  |  |  | 2
3
| 1 | 1 | 01: 7.55; 02: 9.25; 05: 8.6185; 06: < 10; 07: 5.5; 08: 8.7; 09: 9.22; 11: 8.2; 12: 9.825; 13: 9; 14: 9.5 |  |  |  | 1
| 1 | 2 | 01: 7.475; 02: 8.98; 05: 8.6709; 06: 10; 07: 3.9; 08: 8.5; 09: 9.12; 11: 7.6; 12: 10.022; 13: 9; 14: 9.5 |  |  |  | 0
| 1 | 3 | 01: 7.475; 02: 8.47; 05: 8.6203; 06: 11; 07: 3.7; 08: 8.7; 09: 8.77; 11: 8.3; 12: 9.289; 13: 9; 14: 9.5 |  |  |  | 16
| 1 | 4 | 01: 7.34; 02: 8.68; 05: 8.7318; 06: 10; 07: 4; 08: 8.9; 09: 9.36; 11: 8.7; 12: 9.628; 13: 9; 14: 10 |  |  |  | 8.8249
| 1 | 5 | 01: 7.65; 02: 8.46; 05: 8.7387; 06: 10; 07: 4.1; 08: 8.9; 09: 8.73; 11: 10; 12: 9.655; 13: 9; 14: 9.5 |  |  |  | 40
| 1 | 6 | 01: 7.63; 02: 8.3; 05: 8.6043; 06: < 10; 07: 3.8; 08: 8.7; 09: 9.25; 11: 7.9; 12: 9.333; 13: 8; 14: 10
| Mean |  | 01: 7.52; 02: 8.69; 05: 8.6641; 06: 10.25; 07: 4.1667; 08: 8.7333; 09: 9.075; 11: 8.45; 12: 9.6253; 13: 8.8333; 14: 9.6667
| Median |  | 01: 7.5125; 02: 8.575; 05: 8.6456; 06: 10; 07: 3.95; 08: 8.7; 09: 9.17; 11: 8.25; 12: 9.6415; 13: 9; 14: 9.5
| Std Dev. |  | 01: 0.1152; 02: 0.36089; 05: 0.059613; 06: 0.5; 07: 0.66833; 08: 0.15055; 09: 0.26342; 11: 0.84558; 12: 0.28167; 13: 0.40825; 14: 0.2582
| Rel.Std.Dev. |  | 01: 1.532%; 02: 4.153%; 05: 0.688%; 06: 4.878%; 07: 16.04%; 08: 1.724%; 09: 2.903%; 11: 10.01%; 12: 2.926%; 13: 4.622%; 14: 2.671%
| PDM3 |  | 01: -14.79%; 02: -1.529%; 05: -1.823%; 06: 16.15%; 07: -52.79%; 08: -1.038%; 09: 2.834%; 11: -4.249%; 12: 9.07%; 13: 0.09514%; 14: 9.538%
| Z-Score (Absolute) |  | 01: 2.39; 02: 0.09; 05: 0.14; 06: 2.99; 07: 9.54; 08: 0; 09: 0.67; 11: 0.56; 12: 1.76; 13: 0.53; 14: 1.84
| Analytical results for Nd in OREAS 604 (Certified Value 16.5 ppm) | [23 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  |  |  | 1
2
| 1 | 1 | 01: 15.49; 02: 18.1049; 05: 15.9628; 06: 17.67; 07: 16; 08: 17.2 |  |  |  | 1
| 1 | 2 | 01: 15.37; 02: 17.8563; 05: 16.1779; 06: 18.02; 07: 15.2; 08: 15.9 |  |  |  | 0
| 1 | 3 | 01: 15.65; 02: 18.089; 05: 15.811; 06: 17.49; 07: 15.2; 08: 16.6 |  |  |  | 16
| 1 | 4 | 01: 15.42; 02: 17.4304; 05: 16.5903; 06: 18.16; 07: 14.3; 08: 15.9 |  |  |  | 16.5097
| 1 | 5 | 01: 16.17; 02: 17.672; 05: 16.103; 06: 17.34; 07: 16; 08: 16.5 |  |  |  | 41
| 1 | 6 | 01: 15.97; 02: 18.0853; 05: 15.8569; 06: 18.16; 07: 15.5; 08: 15.4
| Mean |  | 01: 15.6783; 02: 17.873; 05: 16.0837; 06: 17.8067; 07: 15.3667; 08: 16.25
| Median |  | 01: 15.57; 02: 17.9708; 05: 16.0329; 06: 17.845; 07: 15.35; 08: 16.2
| Std Dev. |  | 01: 0.324; 02: 0.27646; 05: 0.28507; 06: 0.35551; 07: 0.63456; 08: 0.64109
| Rel.Std.Dev. |  | 01: 2.067%; 02: 1.547%; 05: 1.772%; 06: 1.996%; 07: 4.129%; 08: 3.945%
| PDM3 |  | 01: -5.036%; 02: 8.257%; 05: -2.581%; 06: 7.856%; 07: -6.924%; 08: -1.573%
| Z-Score (Absolute) |  | 01: 0.51; 02: 1.79; 05: 0.09; 06: 1.72; 07: 0.84; 08: 0.09
| Analytical results for Ni in OREAS 604 (Certified Value 638 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 682.81 | 657.2 | 677 | 619 | 665 | 639 | 627.5763 | 761 | 628 | 622.1 | 635.361 | 642 | 683.51 | 614 | 646 | 548 | 616.55 | 532 | 634 | 650 | 645 |  | 1
| 1 | 2 | 685.1882 | 655.1 | 667 | 649 | 658 | 643 | 633.3606 | 795 | 618 | 580.1 | 634.492 | 656 | 629.82 | 614 | 596 | 582 | 601.81 | 534 | 625 | 650 | 640 |  | 0
| 1 | 3 | 685.7422 | 666.2 | 667 | 636 | 661 | 612 | 623.079 | 781 | 622 | 584.7 | 640.37 | 657 | 677.28 | 592 | 618 | 594 | 620.13 | 537 | 636 | 675 | 655 |  | 16
| 1 | 4 | 681.636 | 655.7 | 666 | 629 | 658 | 633 | 621.2522 | 778 | 634 | 599 | 637.604 | 649 | 637.02 | 591 | 626 | 570 | 647.06 | 517 | 655 | 670 | 655 |  | 637.5076
| 1 | 5 | 679.2726 | 669.1 | 651 | 684 | 645 | 631 | 630.1883 | 766 | 634 | 536.1 | 629.889 | 649 | 669.8 | 599 | 681 | 554 | 602.55 | 544 | 629 | 685 | 655 |  | 42
| 1 | 6 | 682.5479 | 662.3 | 657 | 629 | 672 | 641 | 624.5537 | 763 | 627 | 584 | 637.383 | 657 | 640.15 | 610 | 687 | 557 | 619 | 550 | 634 | 690 | 655
| Mean |  | 682.8661 | 660.9333 | 664.1667 | 641 | 659.8333 | 633.1667 | 626.6683 | 774 | 627.1667 | 584.3333 | 635.8498 | 651.6667 | 656.2633 | 603.3333 | 642.3333 | 567.5 | 617.85 | 535.6667 | 635.5 | 670 | 650.8333
| Median |  | 682.6789 | 659.75 | 666.5 | 632.5 | 659.5 | 636 | 626.065 | 772 | 627.5 | 584.35 | 636.372 | 652.5 | 654.975 | 604.5 | 636 | 563.5 | 617.775 | 535.5 | 634 | 672.5 | 655
| Std Dev. |  | 2.3747 | 5.8592 | 9.0425 | 23.2809 | 8.9759 | 11.3563 | 4.5758 | 13.1149 | 6.4005 | 28.2349 | 3.563 | 6.0553 | 23.2229 | 10.6896 | 36.0814 | 17.841 | 16.4487 | 11.3255 | 10.3682 | 17.0294 | 6.6458
| Rel.Std.Dev. |  | 0.3477% | 0.8865% | 1.361% | 3.632% | 1.36% | 1.794% | 0.7302% | 1.694% | 1.021% | 4.832% | 0.5603% | 0.9292% | 3.539% | 1.772% | 5.617% | 3.144% | 2.662% | 2.114% | 1.632% | 2.542% | 1.021%
| PDM3 |  | 7.115% | 3.675% | 4.182% | 0.5478% | 3.502% | -0.6809% | -1.7% | 21.41% | -1.622% | -8.341% | -0.26% | 2.221% | 2.942% | -5.361% | 0.757% | -10.98% | -3.084% | -15.97% | -0.3149% | 5.097% | 2.09%
| Z-Score (Absolute) |  | 1.57 | 0.81 | 0.92 | 0.17 | 0.77 | 0 | 0.37 | 4.71 | 0.35 | 1.5 | 0.05 | 0.49 | 0.65 | 1.18 | 0.17 | 2.41 | 0.67 | 3.51 | 0.2 | 1.12 | 0.46
| Analytical results for P in OREAS 604 (Certified Value 573 ppm) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  |  |  |  | 0
0
| 1 | 1 | 572.6245 | 963 | 570 | 550 | 560 | 540 | 575.7136 | 577 | 681 | 532 | 300 | 539 | 600 | 598.15 | 636 | 470 | 600 | 550 |  |  |  |  | 1
| 1 | 2 | 576.085 | 811 | 560 | 570 | 560 | 540 | 567.4294 | 561 | 631 | 540 | 300 | 519 | 600 | 519.832 | 634 | 480 | 550 | 550 |  |  |  |  | 0
| 1 | 3 | 573.1282 | 833 | 560 | 560 | 560 | 540 | 580.0641 | 564 | 603 | 545 | 300 | 546 | 600 | 649.507 | 654 | 460 | 600 | 550 |  |  |  |  | 16
| 1 | 4 | 570.8467 | 1029 | 560 | 540 | 570 | 570 | 562.1039 | 569 | 634 | 538 | 300 | 526 | 600 | 697.073 | 601 | 490 | 600 | 550 |  |  |  |  | 572.5113
| 1 | 5 | 568.008 | 1008 | 550 | 590 | 540 | 550 | 566.3234 | 571 | 566 | 541 | 200 | 528 | 700 | 671.484 | 640 | 460 | 600 | 550 |  |  |  |  | 43
| 1 | 6 | 576.0677 | 800 | 540 | 560 | 570 | 550 | 593.1259 | 563 | 589 | 544 | 200 | 523 | 600 | 652.09 | 634 | 470 | 650 | 550
| Mean |  | 572.7933 | 907.3333 | 556.6667 | 561.6667 | 560 | 548.3333 | 574.1267 | 567.5 | 617.3333 | 540 | 266.6667 | 530.1667 | 616.6667 | 631.356 | 633.1667 | 471.6667 | 600 | 550
| Median |  | 572.8764 | 898 | 560 | 560 | 560 | 545 | 571.5715 | 566.5 | 617 | 540.5 | 300 | 527 | 600 | 650.7985 | 635 | 470 | 600 | 550
| Std Dev. |  | 3.1108 | 104.2702 | 10.328 | 17.224 | 10.9545 | 11.6905 | 11.3933 | 5.9917 | 40.4013 | 4.6904 | 51.6398 | 10.2648 | 40.8248 | 63.6059 | 17.4633 | 11.6905 | 31.6228 | 0
| Rel.Std.Dev. |  | 0.5431% | 11.49% | 1.855% | 3.067% | 1.956% | 2.132% | 1.984% | 1.056% | 6.544% | 0.8686% | 19.36% | 1.936% | 6.62% | 10.07% | 2.758% | 2.479% | 5.27% | 0%
| PDM3 |  | 0.04926% | 58.48% | -2.768% | -1.894% | -2.185% | -4.223% | 0.2822% | -0.8753% | 7.829% | -5.679% | -53.42% | -7.396% | 7.713% | 10.28% | 10.59% | -17.61% | 4.801% | -3.932%
| Z-Score (Absolute) |  | 0.19 | 7.83 | 0.18 | 0.07 | 0.1 | 0.37 | 0.22 | 0.07 | 1.21 | 0.56 | 6.81 | 0.79 | 1.19 | 2.04 | 1.71 | 2.12 | 0.81 | 0.33
| Analytical results for Pb in OREAS 604 (Certified Value 994 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 1097.0237 | 929.7 | 1075 | 1020 | 1035 | 1040 | 964.5374 | 760 | 925 | 1099.7 | 1008.3628 | 939 | 1022 | 924 | 1047 | 820 | 938.274 | 752 | 979 | 935 | 1090 |  | 1
| 1 | 2 | 1091.575 | 933.1 | 1065 | 1040 | 1020 | 1030 | 960.247 | 570 | 972 | 1031.8 | 1008.842 | 943 | 1024 | 912 | 993 | 871 | 946.341 | 764 | 986 | 950 | 1100 |  | 0
| 1 | 3 | 1093.5786 | 925.5 | 1060 | 1020 | 1040 | 1000 | 974.7858 | 620 | 962 | 1026.3 | 1020.2296 | 918 | 1028 | 879 | 1071 | 869 | 937.551 | 769 | 978 | 949 | 1100 |  | 16
| 1 | 4 | 1108.9686 | 921.7 | 1065 | 1030 | 1025 | 1030 | 950.7206 | 670 | 966 | 1063.2 | 1040.3489 | 944 | 1031 | 872 | 954 | 850 | 955.379 | 742 | 987 | 988 | 1100 |  | 994.348
| 1 | 5 | 1103.7168 | 921.5 | 1050 | 1090 | 1005 | 1025 | 976.1141 | 640 | 967 | 1044.1 | 997.2404 | 970 | 1009 | 915 | 1086 | 792 | 931.977 | 792 | 968 | 938 | 1090 |  | 44
| 1 | 6 | 1088.8659 | 921.5 | 1045 | 1030 | 1035 | 1045 | 982.807 | 690 | 978 | 1062.2 | 1020.0654 | 955 | 1012 | 881 | 1119 | 823 | 952.826 | 793 | 977 | 969 | 1060
| Mean |  | 1097.2881 | 925.5 | 1060 | 1038.3333 | 1026.6667 | 1028.3333 | 968.202 | 658.3333 | 961.6667 | 1054.55 | 1015.8482 | 944.8333 | 1021 | 897.1667 | 1045 | 837.5 | 943.7247 | 768.6667 | 979.1667 | 954.8333 | 1090
| Median |  | 1095.3011 | 923.6 | 1062.5 | 1030 | 1030 | 1030 | 969.6616 | 655 | 966.5 | 1053.15 | 1014.4537 | 943.5 | 1023 | 896.5 | 1059 | 836.5 | 942.3075 | 766.5 | 978.5 | 949.5 | 1095
| Std Dev. |  | 7.6833 | 4.9364 | 10.9545 | 26.3944 | 12.9099 | 15.7056 | 11.8454 | 64.9359 | 18.7901 | 26.8205 | 14.7556 | 17.2907 | 8.7636 | 22.2838 | 61.3547 | 31.1689 | 9.288 | 20.7236 | 6.9113 | 20.1734 | 15.4919
| Rel.Std.Dev. |  | 0.7002% | 0.5334% | 1.033% | 2.542% | 1.257% | 1.527% | 1.223% | 9.864% | 1.954% | 2.543% | 1.453% | 1.83% | 0.8583% | 2.484% | 5.871% | 3.722% | 0.9842% | 2.696% | 0.7058% | 2.113% | 1.421%
| PDM3 |  | 10.35% | -6.924% | 6.603% | 4.424% | 3.25% | 3.418% | -2.629% | -33.79% | -3.287% | 6.054% | 2.162% | -4.98% | 2.68% | -9.773% | 5.094% | -15.77% | -5.091% | -22.7% | -1.527% | -3.974% | 9.62%
| Z-Score (Absolute) |  | 1.62 | 0.74 | 1.11 | 0.67 | 0.65 | 0.67 | 0.15 | 4.4 | 0.14 | 1.03 | 0.5 | 0.47 | 0.57 | 1.12 | 0.9 | 1.94 | 0.49 | 2.89 | 0 | 0.33 | 1.6
| Analytical results for Pr in OREAS 604 (Certified Value 4.55 ppm) | [23 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  | 2
3
| 1 | 1 | 01: 4.16; 02: 4.6007; 05: 4.605; 06: 5.11; 07: 4.76; 08: 4.1 |  | 1
| 1 | 2 | 01: 4.13; 02: 4.6337; 05: 4.7356; 06: 5.05; 07: 4.58; 08: 3.88 |  | 0
| 1 | 3 | 01: 4.25; 02: 4.7021; 05: 4.4814; 06: 4.9; 07: 4.54; 08: 4.1 |  | 16
| 1 | 4 | 01: 4.2; 02: 4.6288; 05: 4.7478; 06: 5.09; 07: 4.72; 08: 4 |  | 4.5488
| 1 | 5 | 01: 4.36; 02: 4.7429; 05: 4.8498; 06: 4.85; 07: 4.97; 08: 4.08 |  | 45
| 1 | 6 | 01: 4.27; 02: 4.7468; 05: 4.5033; 06: 5.21; 07: 4.57; 08: 3.9
| Mean |  | 01: 4.2283; 02: 4.6758; 05: 4.6538; 06: 5.035; 07: 4.69; 08: 4.01
| Median |  | 01: 4.225; 02: 4.6679; 05: 4.6703; 06: 5.07; 07: 4.65; 08: 4.04
| Std Dev. |  | 01: 0.083287; 02: 0.063002; 05: 0.14745; 06: 0.13561; 07: 0.16322; 08: 0.1002
| Rel.Std.Dev. |  | 01: 1.97%; 02: 1.347%; 05: 3.168%; 06: 2.693%; 07: 3.48%; 08: 2.499%
| PDM3 |  | 01: -7.046%; 02: 2.792%; 05: 2.308%; 06: 10.69%; 07: 3.103%; 08: -11.85%
| Z-Score (Absolute) |  | 01: 1.49; 02: 0.04; 05: 0.04; 06: 1.26; 07: 0.09; 08: 2.23
| Analytical results for Rb in OREAS 604 (Certified Value 51 ppm) | [23 empty] | Certified
Rb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 13; 07: 14; 08: 15; 09: 17; 10: 19; 11: 24; 12: 25 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS |  | 0
1
| 1 | 1 | 01: 45.57; 02: 52.389; 05: 54.6; 06: 49.8668; 07: 52; 08: 53.27; 09: 55.1; 10: 44.295; 11: 47.4; 12: 50 |  | 1
| 1 | 2 | 01: 45.99; 02: 51.3093; 05: 50.2; 06: 49.4481; 07: 50.1; 08: 53.59; 09: 52.1; 10: 44.33; 11: 45.4; 12: 52.2 |  | 0
| 1 | 3 | 01: 46.17; 02: 51.7112; 05: 52.1; 06: 49.2091; 07: 51.3; 08: 53.38; 09: 53.6; 10: 43.278; 11: 45.8; 12: 52.2 |  | 16
| 1 | 4 | 01: 46.65; 02: 51.4089; 05: 53.6; 06: 49.3908; 07: 51.5; 08: 54.68; 09: 51.5; 10: 44.526; 11: 47.6; 12: 53.2 |  | 51.0422
| 1 | 5 | 01: 47.44; 02: 51.9714; 05: 53.6; 06: 50.0309; 07: 51.5; 08: 53.95; 09: 55.3; 10: 44.071; 11: 46.8; 12: 53 |  | 46
| 1 | 6 | 01: 47.48; 02: 52.0542; 05: 54.8; 06: 49.8839; 07: 51.6; 08: 51.2; 09: 53.2; 10: 44.011; 11: 48.4; 12: 53.2
| Mean |  | 01: 46.55; 02: 51.8073; 05: 53.15; 06: 49.6383; 07: 51.3333; 08: 53.345; 09: 53.4667; 10: 44.0852; 11: 46.9; 12: 52.3
| Median |  | 01: 46.41; 02: 51.8413; 05: 53.6; 06: 49.6575; 07: 51.5; 08: 53.485; 09: 53.4; 10: 44.183; 11: 47.1; 12: 52.6
| Std Dev. |  | 01: 0.78554; 02: 0.41043; 05: 1.7341; 06: 0.33113; 07: 0.64704; 08: 1.1676; 09: 1.5397; 10: 0.43716; 11: 1.1367; 12: 1.2182
| Rel.Std.Dev. |  | 01: 1.688%; 02: 0.7922%; 05: 3.263%; 06: 0.6671%; 07: 1.26%; 08: 2.189%; 09: 2.88%; 10: 0.9916%; 11: 2.424%; 12: 2.329%
| PDM3 |  | 01: -8.801%; 02: 1.499%; 05: 4.13%; 06: -2.75%; 07: 0.5704%; 08: 4.512%; 09: 4.75%; 10: -13.63%; 11: -8.115%; 12: 2.464%
| Z-Score (Absolute) |  | 01: 1.77; 02: 0.08; 05: 0.56; 06: 0.68; 07: 0.08; 08: 0.78; 09: 0.67; 10: 2.64; 11: 1.65; 12: 0.42
| Analytical results for Re in OREAS 604 (Indicative Value 2 ppb) | [23 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 14; 08: 15; 09: 25 |  | ppb
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS |  | 0
0
| 1 | 1 | 01: < 2; 02: < 2; 05: < 5000; 06: 2; 07: < 50; 08: 10; 09: < 100 |  | 1
| 1 | 2 | 01: 2; 02: 3.116; 05: < 5000; 06: 3; 07: < 50; 08: 10; 09: < 100 |  | 4
| 1 | 3 | 01: 2; 02: 1.9841; 05: < 5000; 06: 3; 07: < 50; 08: 10; 09: < 100 |  | 16
| 1 | 4 | 01: 5; 02: < 2; 05: < 5000; 06: 2; 07: < 50; 08: 10; 09: < 100 |  | 2.0634
| 1 | 5 | 01: < 2; 02: < 2; 05: < 5000; 06: 2; 07: < 50; 08: 20; 09: < 100 |  | 10
| 1 | 6 | 01: < 2; 02: 2.0419; 05: < 5000; 06: 3; 07: < 50; 08: 10; 09: < 100
| Mean |  | 01: 3; 02: 2.3806; 06: 2.5; 08: 11.6667
| Median |  | 01: 2; 02: 2.0419; 06: 2.5; 08: 10
| Std Dev. |  | 01: 1.7321; 02: 0.63747; 06: 0.54772; 08: 4.0825
| Rel.Std.Dev. |  | 01: 57.74%; 02: 26.78%; 06: 21.91%; 08: 34.99%
| PDM3 |  | 01: 45.39%; 02: 15.37%; 06: 21.16%; 08: 465.4%
| Z-Score (Absolute) |  | 01: 0.42; 02: 0.93; 05: NA; 06: 0.42; 07: NA; 08: 12.91; 09: NA
| Analytical results for S in OREAS 604 (Certified Value 4.59 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 23 | 24 | 25 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | PF*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 4.9166 | 4.5928 | 4.78 | 4.56 | 4.67 | 4.54 | 4.6235 | 4.74 | 4.37 | 4.91 | 4.3891 | 4.4308 | 4.455 | N.A. | - | 4.654 | 4.2 | 4.99 | 4.56 |  |  |  | 1
| 1 | 2 | 4.9098 | 4.5248 | 4.75 | 4.71 | 4.6 | 4.57 | 4.413 | 4.77 | 4.27 | 4.64 | 4.358 | 4.482 | 4.472 | N.A. | - | 4.572 | 4.2 | 4.94 | 4.58 |  |  |  | 0
| 1 | 3 | 4.8993 | 4.4738 | 4.7 | 4.64 | 4.62 | 4.53 | 4.3687 | 4.81 | 4.38 | 4.59 | 4.4259 | 4.5093 | 4.523 | N.A. | - | 4.716 | 4.2 | 5.03 | 4.64 |  |  |  | 16
| 1 | 4 | 4.8993 | 4.4159 | 4.73 | 4.62 | 4.62 | 4.58 | 4.6385 | 4.79 | 4.49 | 4.87 | 4.4142 | 4.4554 | 4.508 | N.A. | - | 4.913 | 4.3 | 4.78 | 4.72 |  |  |  | 4.592
| 1 | 5 | 4.914 | 4.4974 | 4.66 | 4.87 | 4.52 | 4.57 | 4.4033 | 4.7 | 4.39 | 4.36 | 4.3818 | 4.4187 | 4.471 | N.A. | - | 4.543 | 4.2 | 5.02 | 4.66 |  |  |  | 47
| 1 | 6 | 4.8993 | 4.4353 | 4.66 | 4.64 | 4.7 | 4.64 | 4.3989 | 4.76 | 4.25 | 4.68 | 4.3621 | 4.4777 | 4.505 | N.A. | - | 4.678 | 4.2 | 4.96 | 4.64
| Mean |  | 4.9064 | 4.49 | 4.7133 | 4.6733 | 4.6217 | 4.5717 | 4.4743 | 4.7617 | 4.3583 | 4.675 | 4.3885 | 4.4623 | 4.489 |  |  | 4.6793 | 4.2167 | 4.9533 | 4.6333
| Median |  | 4.9046 | 4.4856 | 4.715 | 4.64 | 4.62 | 4.57 | 4.4081 | 4.765 | 4.375 | 4.66 | 4.3855 | 4.4665 | 4.4885 |  |  | 4.666 | 4.2 | 4.975 | 4.64
| Std Dev. |  | 0.0080635 | 0.064154 | 0.048854 | 0.10764 | 0.062102 | 0.038687 | 0.12235 | 0.038687 | 0.087731 | 0.20047 | 0.027302 | 0.033985 | 0.026616 |  |  | 0.13168 | 0.040825 | 0.091579 | 0.057504
| Rel.Std.Dev. |  | 0.1643% | 1.429% | 1.036% | 2.303% | 1.344% | 0.8462% | 2.735% | 0.8125% | 2.013% | 4.288% | 0.6221% | 0.7616% | 0.5929% |  |  | 2.814% | 0.9682% | 1.849% | 1.241%
| PDM3 |  | 6.847% | -2.221% | 2.643% | 1.772% | 0.6466% | -0.4422% | -2.562% | 3.695% | -5.088% | 1.808% | -4.43% | -2.824% | -2.242% |  |  | 1.902% | -8.173% | 7.869% | 0.9007%
| Z-Score (Absolute) |  | 1.45 | 0.67 | 0.47 | 0.06 | 0 | 0.25 | 0.75 | 0.71 | 1.34 | 0.27 | 1.18 | 0.81 | 0.67 | NA | NA | 0.29 | 2.06 | 1.86 | 0.06
| Analytical results for Sb in OREAS 604 (Certified Value 167 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 25 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 161.16 | 165.38 | 182 | 169 | 180 | 170 | 161.5009 | 170 | 183 | 117.2514 | 168.3 | 166 | 155 | 164 | 97 | 174.153 | 130 | 156 | 157 |  |  |  | 1
| 1 | 2 | 166.42 | 165.27 | 182 | 177 | 175 | 169 | 161.3604 | 160 | 184 | 123.1023 | 166.6 | 168 | 155 | 159 | 143 | 169.004 | 128 | 159 | 166 |  |  |  | 0
| 1 | 3 | 166.2275 | 158.14 | 177 | 170 | 178 | 169 | 158.774 | 170 | 185 | 120.8959 | 166.6 | 168 | 150 | 162 | 157 | 169.397 | 131 | 163 | 165 |  |  |  | 16
| 1 | 4 | 163.4725 | 162.92 | 181 | 176 | 177 | 172 | 159.892 | 160 | 185 | 118.1544 | 168 | 169 | 146 | 172 | 147 | 173.58 | 126 | 163 | 159 |  |  |  | 166.5979
| 1 | 5 | 161.51 | 162.62 | 174 | 185 | 172 | 167 | 160.514 | 160 | 183 | 121.9786 | 162.1 | 166 | 156 | 191 | 147 | 171.22 | 136 | 162 | 176 |  |  |  | 48
| 1 | 6 | 166.5525 | 157.16 | 174 | 176 | 178 | 171 | 159.4368 | 170 | 183 | 117.948 | 165.3 | 172 | 158 | 167 | 118 | 169.647 | 133 | 162 | 149
| Mean |  | 164.2237 | 161.915 | 178.3333 | 175.5 | 176.6667 | 169.6667 | 160.2463 | 165 | 183.8333 | 119.8885 | 166.15 | 168.1667 | 153.3333 | 169.1667 | 134.8333 | 171.1668 | 130.6667 | 160.8333 | 162
| Median |  | 164.85 | 162.77 | 179 | 176 | 177.5 | 169.5 | 160.203 | 165 | 183.5 | 119.5252 | 166.6 | 168 | 155 | 165.5 | 145 | 170.4335 | 130.5 | 162 | 162
| Std Dev. |  | 2.513 | 3.5108 | 3.8297 | 5.7533 | 2.8048 | 1.7512 | 1.0803 | 5.4772 | 0.98319 | 2.4265 | 2.2616 | 2.2286 | 4.4572 | 11.583 | 22.6664 | 2.23 | 3.559 | 2.7869 | 9.2087
| Rel.Std.Dev. |  | 1.53% | 2.168% | 2.148% | 3.278% | 1.588% | 1.032% | 0.6742% | 3.32% | 0.5348% | 2.024% | 1.361% | 1.325% | 2.907% | 6.847% | 16.81% | 1.303% | 2.724% | 1.733% | 5.684%
| PDM3 |  | -1.425% | -2.811% | 7.044% | 5.343% | 6.044% | 1.842% | -3.813% | -0.9591% | 10.35% | -28.04% | -0.2689% | 0.9416% | -7.962% | 1.542% | -19.07% | 2.742% | -21.57% | -3.46% | -2.76%
| Z-Score (Absolute) |  | 0.08 | 0.4 | 1.88 | 1.48 | 1.64 | 0.67 | 0.63 | 0.03 | 2.64 | 6.22 | 0.19 | 0.47 | 1.59 | 0 | 3.1 | 0.88 | 4.73 | 0.42 | 0.39
| Analytical results for Sc in OREAS 604 (Certified Value 4.81 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 2
2
| 1 | 1 | 4.161 | 4.1 | 5 | 5 | 5 | 5 | 5.0537 | 4 | 4.5 | 5.2204 | 5 | 5 | 4.2 | 6.3 | < 5 | 4.61 | 4 | 6 | 4 |  |  |  | 1
| 1 | 2 | 4.2227 | 4 | 5 | 5 | 5 | 5 | 5.1856 | 5 | 4.2 | 5.1484 | 5 | 5 | 4.2 | 6.1 | < 5 | 4.62 | 5 | 5 | 4 |  |  |  | 0
| 1 | 3 | 4.302 | 3.9 | 5 | 5 | 5 | 4 | 4.9594 | 5 | 4.2 | 5.2935 | 5 | 5 | 3.9 | 6.2 | < 5 | 4.635 | 4 | 6 | 4 |  |  |  | 16
| 1 | 4 | 4.434 | 4 | 5 | 5 | 5 | 5 | 4.975 | 6 | 4.5 | 5.3531 | 5 | 5 | 3.5 | 6.2 | < 5 | 4.85 | 5 | 6 | 4 |  |  |  | 4.8145
| 1 | 5 | 4.118 | 4.2 | 5 | 5 | 5 | 5 | 5.2082 | 5 | 4.2 | 5.2552 | 5 | 5 | 4 | 6.6 | < 5 | 4.654 | 4 | 6 | 4 |  |  |  | 49
| 1 | 6 | 4.2233 | 4.1 | 5 | 5 | 5 | 5 | 5.1564 | 4 | 4.4 | 5.176 | 5 | 5 | 4.2 | 6 | < 5 | 4.447 | 4 | 6 | 4
| Mean |  | 4.2435 | 4.05 | 5 | 5 | 5 | 4.8333 | 5.0897 | 4.8333 | 4.3333 | 5.2411 | 5 | 5 | 4 | 6.2333 |  | 4.636 | 4.3333 | 5.8333 | 4
| Median |  | 4.223 | 4.05 | 5 | 5 | 5 | 5 | 5.1051 | 5 | 4.3 | 5.2378 | 5 | 5 | 4.1 | 6.2 |  | 4.6275 | 4 | 6 | 4
| Std Dev. |  | 0.11235 | 0.10488 | 0 | 0 | 0 | 0.40825 | 0.10874 | 0.75277 | 0.15055 | 0.075834 | 0 | 0 | 0.27568 | 0.20656 |  | 0.12867 | 0.5164 | 0.40825 | 0
| Rel.Std.Dev. |  | 2.648% | 2.59% | 0% | 0% | 0% | 8.447% | 2.136% | 15.57% | 3.474% | 1.447% | 0% | 0% | 6.892% | 3.314% |  | 2.775% | 11.92% | 6.999% | 0%
| PDM3 |  | -11.86% | -15.88% | 3.854% | 3.854% | 3.854% | 0.392% | 5.717% | 0.392% | -9.993% | 8.861% | 3.854% | 3.854% | -16.92% | 29.47% |  | -3.707% | -9.993% | 21.16% | -16.92%
| Z-Score (Absolute) |  | 1.5 | 1.93 | 0.19 | 0.19 | 0.19 | 0.19 | 0.39 | 0.19 | 1.3 | 0.72 | 0.19 | 0.19 | 2.04 | 2.93 | NA | 0.63 | 1.3 | 2.04 | 2.04
| Analytical results for Se in OREAS 604 (Certified Value 43.9 ppm) | [23 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 21 | 24 | 25 | [9 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [9 empty] | 1
1
| 1 | 1 | 49.4648 | 45.5 | 43.8937 | 34 | 47 | 46 | 43.95 | 45 | 64 | 34.824 | 65 | 55 | 40 | [9 empty] | 1
| 1 | 2 | 48.64 | 44.2 | 44.927 | 34 | 45 | 47 | 46.93 | 44 | 49 | 34.732 | 59 | 50 | 40 | [9 empty] | 0
| 1 | 3 | 49.0768 | 45.2 | 44.0356 | 33 | 44 | 45 | 44.04 | 47 | 74 | 33.58 | 68 | 50 | 40 | [9 empty] | 16
| 1 | 4 | 47.9925 | 44.4 | 42.3077 | 36 | 45 | 45 | 46.81 | 48 | 65 | 35.332 | 80 | 55 | 40 | [9 empty] | 43.8515
| 1 | 5 | 48.0303 | 46 | 42.7027 | 33 | 40 | 47 | 45.55 | 50 | 68 | 34.638 | 74 | 55 | 45 | [9 empty] | 50
| 1 | 6 | 49.1468 | 44.2 | 42.4559 | 37 | 44 | 44 | 44.22 | 43 | 61 | 34.421 | 65 | 55 | 45
| Mean |  | 48.7252 | 44.9167 | 43.3871 | 34.5 | 44.1667 | 45.6667 | 45.25 | 46.1667 | 63.5 | 34.5878 | 68.5 | 53.3333 | 41.6667
| Median |  | 48.8584 | 44.8 | 43.2982 | 34 | 44.5 | 45.5 | 44.885 | 46 | 64.5 | 34.685 | 66.5 | 55 | 40
| Std Dev. |  | 0.61244 | 0.76004 | 1.0535 | 1.6432 | 2.3166 | 1.2111 | 1.3828 | 2.6394 | 8.3606 | 0.57904 | 7.4498 | 2.582 | 2.582
| Rel.Std.Dev. |  | 1.257% | 1.692% | 2.428% | 4.763% | 5.245% | 2.652% | 3.056% | 5.717% | 13.17% | 1.674% | 10.88% | 4.841% | 6.197%
| PDM3 |  | 11.11% | 2.429% | -1.059% | -21.33% | 0.7187% | 4.139% | 3.189% | 5.279% | 44.81% | -21.13% | 56.21% | 21.62% | -4.982%
| Z-Score (Absolute) |  | 0.67 | 0.06 | 0.36 | 2.09 | 0.21 | 0.08 | 0 | 0.18 | 3.54 | 2.07 | 4.51 | 1.57 | 0.7
| Analytical results for Sm in OREAS 604 (Certified Value 3.08 ppm) | [23 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 25 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS |  |  |  | 2
3
| 1 | 1 | 01: 3.28; 02: 3.2563; 05: 2.987; 06: 3.2; 07: 2.9; 08: 3.1 |  |  |  | 1
| 1 | 2 | 01: 3.21; 02: 3.3874; 05: 3.001; 06: 3.26; 07: 2.8; 08: 2.8 |  |  |  | 0
| 1 | 3 | 01: 3.29; 02: 3.4467; 05: 2.9433; 06: 3.22; 07: 2.9; 08: 2.95 |  |  |  | 16
| 1 | 4 | 01: 3.24; 02: 3.1856; 05: 2.9651; 06: 3.18; 07: 2.7; 08: 2.9 |  |  |  | 3.0831
| 1 | 5 | 01: 3.28; 02: 3.3376; 05: 2.9632; 06: 3.01; 07: 2.9; 08: 2.9 |  |  |  | 51
| 1 | 6 | 01: 3.28; 02: 3.3078; 05: 2.9821; 06: 3.03; 07: 2.9; 08: 3
| Mean |  | 01: 3.2633; 02: 3.3202; 05: 2.9736; 06: 3.15; 07: 2.85; 08: 2.9417
| Median |  | 01: 3.28; 02: 3.3227; 05: 2.9736; 06: 3.19; 07: 2.9; 08: 2.925
| Std Dev. |  | 01: 0.031411; 02: 0.092927; 05: 0.020472; 06: 0.10431; 07: 0.083666; 08: 0.10206
| Rel.Std.Dev. |  | 01: 0.9626%; 02: 2.799%; 05: 0.6884%; 06: 3.311%; 07: 2.936%; 08: 3.47%
| PDM3 |  | 01: 5.844%; 02: 7.69%; 05: -3.552%; 06: 2.168%; 07: -7.562%; 08: -4.589%
| Z-Score (Absolute) |  | 01: 0.84; 02: 1.08; 05: 0.37; 06: 0.37; 07: 0.89; 08: 0.5
| Analytical results for Sn in OREAS 604 (Certified Value 3.83 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | [8 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [8 empty] | 2
3
| 1 | 1 | < 5 | 4 | 3.9255 | < 10 | 4 | 4 | 3.99 | < 15 | 3.7 | < 20 | 3.414 | 5 | 4.2 | 5 | [8 empty] | 1
| 1 | 2 | < 5 | 3.8 | 3.9761 | < 10 | 3.6 | 4 | 4.26 | < 15 | 3.5 | < 20 | 3.721 | 6 | 4 | 5 | [8 empty] | 0
| 1 | 3 | < 5 | 3.8 | 4.12 | < 10 | 3.7 | 4.1 | 3.97 | < 15 | 3.6 | < 20 | 3.397 | 6 | 4 | 5 | [8 empty] | 16
| 1 | 4 | < 5 | 3.8 | 3.8678 | < 10 | 3.6 | 4.1 | 4.19 | < 15 | 3.6 | < 20 | 3.414 | 5 | 4.4 | 5 | [8 empty] | 3.8331
| 1 | 5 | < 5 | 3.8 | 3.8158 | < 10 | 3.6 | 3.9 | 3.9 | < 15 | 4 | < 20 | 3.388 | 5 | 4 | 5 | [8 empty] | 52
| 1 | 6 | < 5 | 3.8 | 3.9362 | < 10 | 3.8 | 4.1 | 4.04 | < 15 | 3.6 | < 20 | 3.301 | 5 | 4.2 | 5
| Mean |  | 02: 3.8333; 05: 3.9402; 07: 3.7167; 08: 4.0333; 09: 4.0583; 12: 3.6667; 14: 3.4392; 15: 5.3333; 16: 4.1333; 17: 5
| Median |  | 02: 3.8; 05: 3.9309; 07: 3.65; 08: 4.05; 09: 4.015; 12: 3.6; 14: 3.4055; 15: 5; 16: 4.1; 17: 5
| Std Dev. |  | 02: 0.08165; 05: 0.10448; 07: 0.16021; 08: 0.08165; 09: 0.13848; 12: 0.17512; 14: 0.14435; 15: 0.5164; 16: 0.1633; 17: 0
| Rel.Std.Dev. |  | 02: 2.13%; 05: 2.652%; 07: 4.311%; 08: 2.024%; 09: 3.412%; 12: 4.776%; 14: 4.197%; 15: 9.682%; 16: 3.951%; 17: 0%
| PDM3 |  | 02: 0.006406%; 05: 2.795%; 07: -3.037%; 08: 5.224%; 09: 5.876%; 12: -4.342%; 14: -10.28%; 15: 39.14%; 16: 7.833%; 17: 30.44%
| Z-Score (Absolute) |  | NA | 0.55 | 0.14 | NA | 0.8 | 0.14 | 0.21 | NA | 1.14 | NA | 1.78 | 3.97 | 0.43 | 2.99
| Analytical results for Sr in OREAS 604 (Certified Value 398 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | PF*MS |  |  | 0
0
| 1 | 1 | 408.055 | 357.12 | 407 | 352 | 399 | 409 | 419.5719 | 355 | 235 | 435.8661 | 409.3 | 445.02 | 205 | 357 | 345 | 397.87 | 326 | 220 | 437 | 430 |  |  | 1
| 1 | 2 | 468.82 | 367.43 | 400 | 343 | 392 | 426 | 421.481 | 339 | 229.7 | 428.3992 | 398.7 | 415.91 | 181 | 355 | 388 | 414.453 | 338 | 289 | 428 | 430 |  |  | 0
| 1 | 3 | 469.17 | 390.1 | 397 | 351 | 402 | 371 | 422.2227 | 343 | 192.3 | 425.5218 | 424.6 | 438.65 | 186 | 357 | 378 | 384.759 | 347 | 275 | 437 | 430 |  |  | 16
| 1 | 4 | 451.395 | 351.31 | 399 | 356 | 371 | 390 | 421.1353 | 345 | 238.1 | 437.4492 | 410.6 | 431.25 | 196 | 355 | 389 | 387.364 | 333 | 295 | 457 | 430 |  |  | 397.8423
| 1 | 5 | 464.795 | 384.89 | 408 | 362 | 371 | 377 | 428.174 | 349 | 180.2 | 437.4683 | 409.1 | 419.74 | 184 | 360 | 336 | 392.445 | 359 | 297 | 473 | 440 |  |  | 53
| 1 | 6 | 465.02 | 371.14 | 389 | 362 | 424 | 392 | 427.4669 | 331 | 246.1 | 430.6391 | 428.2 | 420.14 | 166 | 362 | 372 | 400.385 | 365 | 209 | 445 | 440
| Mean |  | 454.5425 | 370.3317 | 400 | 354.3333 | 393.1667 | 394.1667 | 423.342 | 343.6667 | 220.2333 | 432.5573 | 413.4167 | 428.4517 | 186.3333 | 357.6667 | 368 | 396.2127 | 344.6667 | 264.1667 | 446.1667 | 433.3333
| Median |  | 464.9075 | 369.285 | 399.5 | 354 | 395.5 | 391 | 421.8518 | 344 | 232.35 | 433.2526 | 409.95 | 425.695 | 185 | 357 | 375 | 395.1575 | 342.5 | 282 | 441 | 430
| Std Dev. |  | 23.6799 | 15.1537 | 6.9857 | 7.2847 | 20.2328 | 20.4296 | 3.5822 | 8.2624 | 27.1253 | 5.0885 | 10.9883 | 11.7356 | 13.3367 | 2.8048 | 22.4054 | 10.7382 | 15.1877 | 39.3874 | 16.3514 | 5.164
| Rel.Std.Dev. |  | 5.21% | 4.092% | 1.746% | 2.056% | 5.146% | 5.183% | 0.8462% | 2.404% | 12.32% | 1.176% | 2.658% | 2.739% | 7.157% | 0.7842% | 6.088% | 2.71% | 4.406% | 14.91% | 3.665% | 1.192%
| PDM3 |  | 14.25% | -6.915% | 0.5424% | -10.94% | -1.175% | -0.9239% | 6.409% | -13.62% | -44.64% | 8.726% | 3.915% | 7.694% | -53.16% | -10.1% | -7.501% | -0.4096% | -13.37% | -33.6% | 12.15% | 8.921%
| Z-Score (Absolute) |  | 1.26 | 0.42 | 0.11 | 0.71 | 0.01 | 0.01 | 0.53 | 0.9 | 3.12 | 0.7 | 0.36 | 0.63 | 3.73 | 0.65 | 0.46 | 0.05 | 0.88 | 2.33 | 0.95 | 0.71
| Analytical results for Ta in OREAS 604 (Certified Value < 1 ppm) | [23 empty] | Certified
Ta | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 12; 06: 14; 07: 15; 08: 17; 09: 19; 11: 24; 12: 25 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS |  |  | 2
2
| 1 | 1 | 01: 1.01; 02: 0.62827; 05: 0.23; 06: 0.68; 07: 0.82; 08: 0.31; 09: 0.941; 11: 0.5; 12: 0.6 |  |  | 1
| 1 | 2 | 01: 1.04; 02: 0.60888; 05: 0.22; 06: 0.68; 07: 0.91; 08: 0.36; 09: 0.928; 11: 0.6; 12: 0.6 |  |  | 6
| 1 | 3 | 01: 0.94; 02: 0.6103; 05: 0.19; 06: 0.65; 07: 0.77; 08: 0.46; 09: 0.86; 11: 0.7; 12: 0.6 |  |  | 16
| 1 | 4 | 01: 1; 02: 0.60042; 05: 0.21; 06: 0.67; 07: 0.92; 08: 0.7; 09: 0.919; 11: 0.7; 12: 0.6 |  |  | < 1
| 1 | 5 | 01: 1.18; 02: 0.61897; 05: 0.18; 06: 0.73; 07: 0.86; 08: 1.39; 09: 0.939; 11: 0.4; 12: 0.6 |  |  | 54
| 1 | 6 | 01: 1; 02: 0.6198; 05: 0.22; 06: 0.69; 07: 0.85; 08: 0.41; 09: 0.888; 11: 0.6; 12: 0.6
| Mean |  | 01: 1.0283; 02: 0.61444; 05: 0.20833; 06: 0.68333; 07: 0.855; 08: 0.605; 09: 0.9125; 11: 0.58333; 12: 0.6
| Median |  | 01: 1.005; 02: 0.61464; 05: 0.215; 06: 0.68; 07: 0.855; 08: 0.435; 09: 0.9235; 11: 0.6; 12: 0.6
| Std Dev. |  | 01: 0.081097; 02: 0.0098504; 05: 0.019408; 06: 0.026583; 07: 0.056125; 08: 0.40776; 09: 0.032105; 11: 0.1169; 12: 0
| Rel.Std.Dev. |  | 01: 7.886%; 02: 1.603%; 05: 9.316%; 06: 3.89%; 07: 6.564%; 08: 67.4%; 09: 3.518%; 11: 20.04%; 12: 0%
PDM3
| Z-Score (Absolute) |  | 01: 1.55; 02: 0; 05: 1.65; 06: 0.28; 07: 0.97; 08: 0.67; 09: 1.21; 11: 0.13; 12: 0.06
| Analytical results for Tb in OREAS 604 (Certified Value 0.33 ppm) | [23 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 13; 06: 15; 07: 17; 08: 19; 09: 25 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*OES/MS; 09: 4A*MS |  |  | 2
3
| 1 | 1 | 01: 0.31; 02: 0.35737; 05: 0.31349; 06: 0.37; 07: 0.36; 08: 0.313; 09: 0.32 |  |  | 1
| 1 | 2 | 01: 0.31; 02: 0.37371; 05: 0.32852; 06: 0.34; 07: 0.34; 08: 0.312; 09: 0.26 |  |  | 0
| 1 | 3 | 01: 0.32; 02: 0.35166; 05: 0.28933; 06: 0.34; 07: 0.35; 08: 0.296; 09: 0.28 |  |  | 16
| 1 | 4 | 01: 0.32; 02: 0.35622; 05: 0.31863; 06: 0.34; 07: 0.34; 08: 0.314; 09: 0.28 |  |  | 0.32573
| 1 | 5 | 01: 0.32; 02: 0.3825; 05: 0.3111; 06: 0.33; 07: 0.37; 08: 0.312; 09: 0.28 |  |  | 55
| 1 | 6 | 01: 0.33; 02: 0.37668; 05: 0.31546; 06: 0.36; 07: 0.34; 08: 0.313; 09: 0.28
| Mean |  | 01: 0.31833; 02: 0.36636; 05: 0.31276; 06: 0.34667; 07: 0.35; 08: 0.31; 09: 0.28333
| Median |  | 01: 0.32; 02: 0.36554; 05: 0.31448; 06: 0.34; 07: 0.345; 08: 0.3125; 09: 0.28
| Std Dev. |  | 01: 0.0075277; 02: 0.012814; 05: 0.012979; 06: 0.015055; 07: 0.012649; 08: 0.0068993; 09: 0.019664
| Rel.Std.Dev. |  | 01: 2.365%; 02: 3.498%; 05: 4.15%; 06: 4.343%; 07: 3.614%; 08: 2.226%; 09: 6.94%
| PDM3 |  | 01: -2.271%; 02: 12.47%; 05: -3.983%; 06: 6.428%; 07: 7.451%; 08: -4.829%; 09: -13.02%
| Z-Score (Absolute) |  | 01: 0; 02: 1.14; 05: 0.13; 06: 0.67; 07: 0.75; 08: 0.2; 09: 1.01
| Analytical results for Te in OREAS 604 (Certified Value 25.1 ppm) | [23 empty] | Certified
Te | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 11: Lab; 12: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 11; 06: 12; 07: 13; 08: 14; 09: 15; 11: 17; 12: 25 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*OES; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES; 11: 4A*OES/MS; 12: 4A*MS |  |  | 1
1
| 1 | 1 | 01: 31.3; 02: 24.6032; 05: 22; 06: 9.19; 07: 11.3812; 08: 27.7; 09: 27; 11: 19.4; 12: 23.2 |  |  | 1
| 1 | 2 | 01: 30.5; 02: 24.6087; 05: 22; 06: 9.24; 07: 11.8531; 08: 27; 09: 30; 11: 18; 12: 24.8 |  |  | 0
| 1 | 3 | 01: 29.8; 02: 24.4108; 05: 26; 06: 8.18; 07: 12.089; 08: 26.2; 09: 31; 11: 18.4; 12: 23.8 |  |  | 16
| 1 | 4 | 01: 29.3; 02: 24.0273; 05: 24; 06: 10.76; 07: 11.588; 08: 26.4; 09: 28; 11: 18.4; 12: 24.2 |  |  | 25.1077
| 1 | 5 | 01: 29.4; 02: 24.5648; 05: 25; 06: 9.9; 07: 12.0113; 08: 25.2; 09: 26; 11: 19.7; 12: 23.8 |  |  | 56
| 1 | 6 | 01: 29.3; 02: 24.2094; 05: 22; 06: 9.22; 07: 11.7991; 08: 26.4; 09: 32; 11: 18.1; 12: 22.8
| Mean |  | 01: 29.9333; 02: 24.404; 05: 23.5; 06: 9.415; 07: 11.787; 08: 26.4833; 09: 29; 11: 18.6667; 12: 23.7667
| Median |  | 01: 29.6; 02: 24.4878; 05: 23; 06: 9.23; 07: 11.8261; 08: 26.4; 09: 29; 11: 18.4; 12: 23.8
| Std Dev. |  | 01: 0.81158; 02: 0.2397; 05: 1.7607; 06: 0.85892; 07: 0.26471; 08: 0.83526; 09: 2.3664; 11: 0.70899; 12: 0.70899
| Rel.Std.Dev. |  | 01: 2.711%; 02: 0.9822%; 05: 7.492%; 06: 9.123%; 07: 2.246%; 08: 3.154%; 09: 8.16%; 11: 3.798%; 12: 2.983%
| PDM3 |  | 01: 19.22%; 02: -2.803%; 05: -6.403%; 06: -62.5%; 07: -53.05%; 08: 5.479%; 09: 15.5%; 11: -25.65%; 12: -5.341%
| Z-Score (Absolute) |  | 01: 0.82; 02: 0.08; 05: 0.04; 06: 1.9; 07: 1.58; 08: 0.36; 09: 0.69; 11: 0.67; 12: 0
| Analytical results for Th in OREAS 604 (Certified Value 7.47 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [6 empty] | 2
2
| 1 | 1 | 6.4 | < 20 | < 20 | < 20 | < 20 | 7.6636 | 2.2 | 8.8245 | 7.24 | 8.75 | 5.7 | < 20 | 7.33 | 4 | 6.9 | 7 | [6 empty] | 1
| 1 | 2 | 6.38 | < 20 | < 20 | < 20 | < 20 | 7.3977 | 2.3 | 8.5822 | 6.92 | 8.73 | 5.7 | < 20 | 7.31 | 6 | 8 | 7.2 | [6 empty] | 0
| 1 | 3 | 6.78 | < 20 | < 20 | < 20 | < 20 | 7.7706 | 1.9 | 8.6248 | 7.37 | 8.15 | 5.7 | < 20 | 6.12 | 6 | 8.6 | 7.6 | [6 empty] | 16
| 1 | 4 | 6.49 | < 20 | < 20 | < 20 | < 20 | 7.498 | 2.4 | 8.9777 | 7.43 | 8.29 | 7.2 | < 20 | 6.73 | 5 | 8.8 | 7.7 | [6 empty] | 7.4708
| 1 | 5 | 7.09 | < 20 | < 20 | < 20 | < 20 | 7.7768 | 1.5 | 8.7922 | 8.05 | 7.74 | 7.7 | < 20 | 6.88 | 4 | 7.9 | 7.7 | [6 empty] | 57
| 1 | 6 | 6.9 | < 20 | < 20 | < 20 | < 20 | 7.9169 | 2.4 | 8.6394 | 8.01 | 8.44 | 5.6 | < 20 | 6.45 | 3 | 8.2 | 7.5
| Mean |  | 6.6733 |  |  |  |  | 7.6706 | 2.1167 | 8.7401 | 7.5033 | 8.35 | 6.2667 |  | 6.8033 | 4.6667 | 8.0667 | 7.45
| Median |  | 6.635 |  |  |  |  | 7.7171 | 2.25 | 8.7158 | 7.4 | 8.365 | 5.7 |  | 6.805 | 4.5 | 8.1 | 7.55
| Std Dev. |  | 0.29351 |  |  |  |  | 0.19299 | 0.35449 | 0.15141 | 0.4446 | 0.38163 | 0.93095 |  | 0.47689 | 1.2111 | 0.66833 | 0.2881
| Rel.Std.Dev. |  | 4.398% |  |  |  |  | 2.516% | 16.75% | 1.732% | 5.925% | 4.57% | 14.86% |  | 7.01% | 25.95% | 8.285% | 3.867%
| PDM3 |  | -10.67% |  |  |  |  | 2.674% | -71.67% | 16.99% | 0.4352% | 11.77% | -16.12% |  | -8.935% | -37.53% | 7.976% | -0.2787%
| Z-Score (Absolute) |  | 0.67 | NA | NA | NA | NA | 0.19 | 4.52 | 1.12 | 0.05 | 0.78 | 1.28 | NA | 0.56 | 2.42 | 0.54 | 0
| Analytical results for Ti in OREAS 604 (Certified Value 0.191 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | PF*OES |  |  | 3
3
| 1 | 1 | 0.174 | 0.191 | 0.2 | 0.19 | 0.2 | 0.19 | 0.20466 | 0.18 | 0.173 | 0.19635 | 0.1909 | 0.19 | 0.15 | 0.17 | 0.2 | 0.17 | 0.1315 | 0.18 | 0.215 | 0.21 |  |  | 1
| 1 | 2 | 0.185 | 0.1901 | 0.2 | 0.2 | 0.2 | 0.18 | 0.20741 | 0.16 | 0.157 | 0.1995 | 0.1909 | 0.19 | 0.15 | 0.18 | 0.2 | 0.172 | 0.1327 | 0.19 | 0.21 | 0.21 |  |  | 0
| 1 | 3 | 0.18 | 0.1903 | 0.2 | 0.2 | 0.2 | 0.18 | 0.20629 | 0.17 | 0.154 | 0.19915 | 0.1881 | 0.19 | 0.14 | 0.18 | 0.2 | 0.17 | 0.1352 | 0.19 | 0.21 | 0.2 |  |  | 16
| 1 | 4 | 0.18 | 0.1932 | 0.2 | 0.2 | 0.2 | 0.19 | 0.20657 | 0.17 | 0.16 | 0.20137 | 0.193 | 0.19 | 0.11 | 0.18 | 0.2 | 0.173 | 0.1308 | 0.19 | 0.215 | 0.21 |  |  | 0.19057
| 1 | 5 | 0.18 | 0.1958 | 0.19 | 0.21 | 0.19 | 0.19 | 0.20802 | 0.17 | 0.158 | 0.19389 | 0.1992 | 0.19 | 0.14 | 0.18 | 0.2 | 0.169 | 0.1364 | 0.19 | 0.215 | 0.21 |  |  | 58
| 1 | 6 | 0.185 | 0.1913 | 0.19 | 0.2 | 0.2 | 0.19 | 0.20822 | 0.17 | 0.159 | 0.19787 | 0.196 | 0.19 | 0.14 | 0.18 | 0.19 | 0.169 | 0.1411 | 0.18 | 0.23 | 0.21
| Mean |  | 0.18067 | 0.19195 | 0.19667 | 0.2 | 0.19833 | 0.18667 | 0.20686 | 0.17 | 0.16017 | 0.19802 | 0.19302 | 0.19 | 0.13833 | 0.17833 | 0.19833 | 0.1705 | 0.13462 | 0.18667 | 0.21583 | 0.20833
| Median |  | 0.18 | 0.19115 | 0.2 | 0.2 | 0.2 | 0.19 | 0.20699 | 0.17 | 0.1585 | 0.19851 | 0.19195 | 0.19 | 0.14 | 0.18 | 0.2 | 0.17 | 0.13395 | 0.19 | 0.215 | 0.21
| Std Dev. |  | 0.0040825 | 0.0021843 | 0.005164 | 0.0063246 | 0.0040825 | 0.005164 | 0.0013219 | 0.0063246 | 0.0066156 | 0.0026305 | 0.0040067 | 3.0405e-17 | 0.01472 | 0.0040825 | 0.0040825 | 0.0016432 | 0.0038343 | 0.005164 | 0.0073598 | 0.0040825
| Rel.Std.Dev. |  | 2.26% | 1.138% | 2.626% | 3.162% | 2.058% | 2.766% | 0.639% | 3.72% | 4.13% | 1.328% | 2.076% | 1.6e-14% | 10.64% | 2.289% | 2.058% | 0.9637% | 2.848% | 2.766% | 3.41% | 1.96%
| PDM3 |  | -5.199% | 0.7215% | 3.196% | 4.946% | 4.071% | -2.051% | 8.546% | -10.8% | -15.96% | 3.907% | 1.281% | -0.3017% | -27.41% | -6.424% | 4.071% | -10.53% | -29.36% | -2.051% | 13.25% | 9.318%
| Z-Score (Absolute) |  | 0.72 | 0.07 | 0.4 | 0.63 | 0.51 | 0.3 | 1.11 | 1.46 | 2.15 | 0.49 | 0.14 | 0.07 | 3.28 | 0.88 | 0.51 | 1.43 | 3.93 | 0.3 | 1.73 | 1.21
| Analytical results for Tl in OREAS 604 (Certified Value 7.44 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 24 | 25 | [8 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [8 empty] | 2
2
| 1 | 1 | 5.69 | 8.34 | 7.4299 | < 5 | 8.16 | 6.2318 | 7.61 | 8.45 | 8.23 | < 20 | 7.898 | < 2 | 7.6 | 7.4 | [8 empty] | 1
| 1 | 2 | 6.3267 | 8.54 | 7.5043 | 7 | 7.48 | 6.204 | 7.82 | 8.27 | 8.03 | < 20 | 7.612 | < 2 | 7.7 | 7.6 | [8 empty] | 0
| 1 | 3 | 6.68 | 8.3 | 7.5059 | 5 | 7.99 | 6.2347 | 7.51 | 7.82 | 8.34 | < 20 | 7.592 | < 2 | 7.7 | 7.5 | [8 empty] | 16
| 1 | 4 | 4.9767 | 8.47 | 7.3798 | 8 | 7.79 | 6.3998 | 7.71 | 8.07 | 8.05 | < 20 | 7.855 | < 2 | 7.8 | 7.6 | [8 empty] | 7.4435
| 1 | 5 | 6.19 | 8.54 | 7.4916 | 6 | 8.03 | 6.4605 | 7.73 | 7.52 | 8.41 | < 20 | 7.663 | 3 | 7.7 | 7.4 | [8 empty] | 59
| 1 | 6 | 6.1233 | 8.23 | 7.3903 | 6 | 8.19 | 6.2923 | 7.72 | 8.18 | 7.76 | < 20 | 7.713 | < 2 | 7.9 | 7.5
| Mean |  | 5.9978 | 8.4033 | 7.4503 | 6.4 | 7.94 | 6.3038 | 7.6833 | 8.0517 | 8.1367 |  | 7.7222 | 3 | 7.7333 | 7.5
| Median |  | 6.1567 | 8.405 | 7.4608 | 6 | 8.01 | 6.2635 | 7.715 | 8.125 | 8.14 |  | 7.688 | 3 | 7.7 | 7.5
| Std Dev. |  | 0.59392 | 0.13155 | 0.057798 | 1.1402 | 0.26668 | 0.10374 | 0.10801 | 0.33451 | 0.23897 |  | 0.12746 |  | 0.10328 | 0.089443
| Rel.Std.Dev. |  | 9.902% | 1.566% | 0.7758% | 17.82% | 3.359% | 1.646% | 1.406% | 4.155% | 2.937% |  | 1.651% |  | 1.336% | 1.193%
| PDM3 |  | -19.42% | 12.89% | 0.09089% | -14.02% | 6.67% | -15.31% | 3.222% | 8.17% | 9.312% |  | 3.743% | -59.7% | 3.893% | 0.7586%
| Z-Score (Absolute) |  | 3.09 | 1.32 | 0.43 | 2.35 | 0.47 | 2.53 | 0 | 0.67 | 0.83 | NA | 0.07 | 8.57 | 0.09 | 0.34
| Analytical results for Tm in OREAS 604 (Certified Value 81 ppb) | [23 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab; 07: Lab |  |  | 1
| No | No | 01: 02; 02: 09; 05: 15; 06: 17; 07: 25 |  |  | ppb
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS |  |  | 0
0
| 1 | 1 | 01: 80; 02: 91.3358; 05: 100; 06: 70; 07: 100 |  |  | 1
| 1 | 2 | 01: 80; 02: 86.8003; 05: 100; 06: 70; 07: 80 |  |  | 7
| 1 | 3 | 01: 80; 02: 85.3795; 05: 100; 06: 70; 07: 80 |  |  | 16
| 1 | 4 | 01: 80; 02: 82.975; 05: < 100; 06: 70; 07: 60 |  |  | 81.1841
| 1 | 5 | 01: 90; 02: 90.903; 05: < 100; 06: 80; 07: 80 |  |  | 60
| 1 | 6 | 01: 80; 02: 98.1284; 05: 100; 06: 70; 07: 80
| Mean |  | 01: 81.6667; 02: 89.2537; 05: 100; 06: 71.6667; 07: 80
| Median |  | 01: 80; 02: 88.8516; 05: 100; 06: 70; 07: 80
| Std Dev. |  | 01: 4.0825; 02: 5.4086; 05: 0; 06: 4.0825; 07: 12.6491
| Rel.Std.Dev. |  | 01: 4.999%; 02: 6.06%; 05: 0%; 06: 5.696%; 07: 15.81%
| PDM3 |  | 01: 0.5945%; 02: 9.94%; 05: 23.18%; 06: -11.72%; 07: -1.458%
| Z-Score (Absolute) |  | 01: 0; 02: 3.07; 05: 0.67; 06: 4.05; 07: 0.67
| Analytical results for U in OREAS 604 (Certified Value 3.17 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [7 empty] | 2
3
| 1 | 1 | 3.04 | < 10 | < 10 | 2.9661 | < 10 | 3.4 | 3.3018 | 3.1 | 3.41 | 2.68 | < 20 | 3.06 | < 20 | 3.2 | 3 | [7 empty] | 1
| 1 | 2 | 3.06 | < 10 | < 10 | 3.1004 | < 10 | 3.2 | 3.3783 | 3.13 | 3.3 | 2.53 | < 20 | 3.06 | < 20 | 3.2 | 3 | [7 empty] | 0
| 1 | 3 | 3.14 | < 10 | < 10 | 3.0484 | < 10 | 3.4 | 3.4001 | 3.03 | 3.12 | 2.58 | < 20 | 2.91 | < 20 | 3.2 | 3.1 | [7 empty] | 16
| 1 | 4 | 3.13 | < 10 | < 10 | 3.014 | < 10 | 3.2 | 3.3253 | 3.13 | 3.26 | 3.17 | < 20 | 3.17 | < 20 | 3.5 | 3.1 | [7 empty] | 3.1653
| 1 | 5 | 3.17 | < 10 | < 10 | 2.952 | < 10 | 3.3 | 3.3231 | 3.18 | 3.04 | 2.61 | < 20 | 3.12 | < 20 | 3.3 | 3.1 | [7 empty] | 61
| 1 | 6 | 3.12 | < 10 | < 10 | 2.9816 | < 10 | 3.4 | 3.3848 | 3.09 | 3.23 | 2.5 | < 20 | 3.05 | < 20 | 3 | 3.1
| Mean |  | 3.11 |  |  | 3.0104 |  | 3.3167 | 3.3522 | 3.11 | 3.2267 | 2.6783 |  | 3.0617 |  | 3.2333 | 3.0667
| Median |  | 3.125 |  |  | 2.9978 |  | 3.35 | 3.3518 | 3.115 | 3.245 | 2.595 |  | 3.06 |  | 3.2 | 3.1
| Std Dev. |  | 0.0498 |  |  | 0.056154 |  | 0.098319 | 0.040385 | 0.0502 | 0.1314 | 0.24895 |  | 0.087502 |  | 0.1633 | 0.05164
| Rel.Std.Dev. |  | 1.601% |  |  | 1.865% |  | 2.964% | 1.205% | 1.614% | 4.072% | 9.295% |  | 2.858% |  | 5.05% | 1.684%
| PDM3 |  | -1.747% |  |  | -4.893% |  | 4.782% | 5.905% | -1.747% | 1.939% | -15.38% |  | -3.274% |  | 2.15% | -3.116%
| Z-Score (Absolute) |  | 0 | NA | NA | 0.62 | NA | 1.29 | 1.51 | 0 | 0.73 | 3.31 | NA | 0.3 | NA | 0.77 | 0.27
| Analytical results for V in OREAS 604 (Certified Value 36 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  |  | 1
2
| 1 | 1 | 36.1077 | 34 | 35 | 36 | 36 | 34 | 38.9076 | 39 | 31 | 35.3236 | 37 | 38 | 35 | 40 | 38 | 34.81 | 34 | 35 | 40 | 35 |  |  | 1
| 1 | 2 | 37.2107 | 34 | 36 | 37 | 35 | 35 | 38.7631 | 38 | 29 | 36.6413 | 37 | 38 | 34 | 40 | 37 | 34.752 | 34 | 35 | 40 | 35 |  |  | 0
| 1 | 3 | 36.8795 | 34 | 34 | 37 | 36 | 33 | 38.2346 | 39 | 28 | 35.8708 | 35 | 38 | 32 | 42 | 37 | 34.863 | 35 | 35 | 40 | 35 |  |  | 16
| 1 | 4 | 36.9108 | 34 | 35 | 37 | 35 | 34 | 37.9834 | 40 | 30 | 36.3243 | 37 | 38 | 32 | 40 | 36 | 35.727 | 33 | 36 | 40 | 35 |  |  | 36.0056
| 1 | 5 | 36.7049 | 34 | 34 | 40 | 35 | 34 | 36.9009 | 40 | 28 | 34.9128 | 36 | 38 | 35 | 41 | 35 | 34.169 | 35 | 35 | 40 | 35 |  |  | 62
| 1 | 6 | 37.4723 | 34 | 35 | 37 | 36 | 35 | 36.9807 | 37 | 29 | 35.0386 | 35 | 39 | 34 | 40 | 34 | 34.313 | 35 | 35 | 40 | 35
| Mean |  | 36.881 | 34 | 34.8333 | 37.3333 | 35.5 | 34.1667 | 37.9617 | 38.8333 | 29.1667 | 35.6852 | 36.1667 | 38.1667 | 33.6667 | 40.5 | 36.1667 | 34.7723 | 34.3333 | 35.1667 | 40 | 35
| Median |  | 36.8951 | 34 | 35 | 37 | 35.5 | 34 | 38.109 | 39 | 29 | 35.5972 | 36.5 | 38 | 34 | 40 | 36.5 | 34.781 | 34.5 | 35 | 40 | 35
| Std Dev. |  | 0.46651 | 0 | 0.75277 | 1.3663 | 0.54772 | 0.75277 | 0.86017 | 1.169 | 1.169 | 0.70746 | 0.98319 | 0.40825 | 1.3663 | 0.83666 | 1.472 | 0.54707 | 0.8165 | 0.40825 | 0 | 0
| Rel.Std.Dev. |  | 1.265% | 0% | 2.161% | 3.66% | 1.543% | 2.203% | 2.266% | 3.01% | 4.008% | 1.982% | 2.719% | 1.07% | 4.058% | 2.066% | 4.07% | 1.573% | 2.378% | 1.161% | 0% | 0%
| PDM3 |  | 2.431% | -5.57% | -3.256% | 3.688% | -1.404% | -5.107% | 5.433% | 7.854% | -18.99% | -0.8897% | 0.4474% | 6.002% | -6.496% | 12.48% | 0.4474% | -3.425% | -4.644% | -2.33% | 11.09% | -2.793%
| Z-Score (Absolute) |  | 0.68 | 0.84 | 0.4 | 0.64 | 0.05 | 0.75 | 1.25 | 1.72 | 3.4 | 0.05 | 0.3 | 1.36 | 1.02 | 2.6 | 0.3 | 0.43 | 0.67 | 0.23 | 2.33 | 0.31
| Analytical results for W in OREAS 604 (Certified Value 16.7 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [7 empty] | 1
1
| 1 | 1 | 11.4233 | 17.9 | 16.6896 | 15 | 22.9 | 14.7112 | 17.3 | 16.84 | 39 | 25.1 | < 20 | 19.42 | 16 | 18 | 16 | [7 empty] | 1
| 1 | 2 | 12.3267 | 18.6 | 16.6444 | 13 | 22.4 | 15.1915 | 17.6 | 15.32 | 38 | 26 | < 20 | 19.27 | 14 | 18.5 | 17.5 | [7 empty] | 0
| 1 | 3 | 11.67 | 17.9 | 16.844 | 14 | 21.9 | 14.9047 | 17.4 | 16.49 | 35 | 27.1 | < 20 | 18.09 | 14 | 19 | 17 | [7 empty] | 16
| 1 | 4 | 12.36 | 19.5 | 16.2436 | 14 | 23.9 | 14.9192 | 17 | 15.31 | 35 | 27.5 | < 20 | 19 | 14 | 19.5 | 17 | [7 empty] | 16.7237
| 1 | 5 | 11.8567 | 18.9 | 16.2274 | 14 | 23.3 | 15.1532 | 17 | 16.58 | 39 | 32.1 | < 20 | 18.74 | 13 | 16 | 17 | [7 empty] | 63
| 1 | 6 | 11.5467 | 17.8 | 16.6362 | 13 | 22.9 | 14.6405 | 17.4 | 15.64 | 36 | 25.2 | < 20 | 18.02 | 15 | 18.5 | 17
| Mean |  | 11.8639 | 18.4333 | 16.5476 | 13.8333 | 22.8833 | 14.92 | 17.2833 | 16.03 | 37 | 27.1667 |  | 18.7567 | 14.3333 | 18.25 | 16.9167
| Median |  | 11.7633 | 18.25 | 16.6403 | 14 | 22.9 | 14.912 | 17.35 | 16.065 | 37 | 26.55 |  | 18.87 | 14 | 18.5 | 17
| Std Dev. |  | 0.39816 | 0.68605 | 0.25306 | 0.75277 | 0.69402 | 0.22365 | 0.24014 | 0.68481 | 1.8974 | 2.6059 |  | 0.59163 | 1.0328 | 1.2145 | 0.4916
| Rel.Std.Dev. |  | 3.356% | 3.722% | 1.529% | 5.442% | 3.033% | 1.499% | 1.389% | 4.272% | 5.128% | 9.592% |  | 3.154% | 7.206% | 6.655% | 2.906%
| PDM3 |  | -29.06% | 10.22% | -1.053% | -17.28% | 36.83% | -10.79% | 3.346% | -4.148% | 121.2% | 62.44% |  | 12.16% | -14.29% | 9.126% | 1.154%
| Z-Score (Absolute) |  | 1.87 | 0.44 | 0.23 | 1.18 | 2 | 0.8 | 0.03 | 0.41 | 6.96 | 3.16 | NA | 0.55 | 1 | 0.53 | 0.03
| Analytical results for Y in OREAS 604 (Certified Value 7.16 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 7.1714 | 6.66 | 7.4357 | 7 | 7.4 | 7.5185 | 7.4 | 7 | < 10 | 7.5 | 6 | 6.94 | 15 | 5 | 7.4 | 7.1 | [6 empty] | 1
| 1 | 2 | 6.9777 | 6.39 | 7.2813 | 7 | 6.8 | 8.1164 | 6.9 | 7 | < 10 | 7.2 | 6 | 6.96 | 16 | 6 | 6.9 | 7.4 | [6 empty] | 0
| 1 | 3 | 7.0945 | 6.49 | 7.1366 | 7 | 7.2 | 7.4186 | 7.2 | 7 | < 10 | 7.4 | 6 | 6.83 | 15 | 5 | 7.1 | 7.2 | [6 empty] | 16
| 1 | 4 | 7.2797 | 6.64 | 7.3288 | 7 | 7.3 | 7.8302 | 7.2 | 7 | < 10 | 7.3 | 6 | 7.06 | 15 | 6 | 7.5 | 7.4 | [6 empty] | 7.1644
| 1 | 5 | 6.9953 | 6.7 | 7.2735 | 7 | 7.2 | 7.9295 | 7.1 | 7 | < 10 | 7.8 | 6 | 7.02 | 16 | 6 | 7.2 | 7.4 | [6 empty] | 64
| 1 | 6 | 7.08 | 6.62 | 7.4114 | 7 | 7.5 | 7.4112 | 7.1 | 7 | < 10 | 7.1 | 6 | 6.87 | 15 | 5 | 7.6 | 7.4
| Mean |  | 7.0998 | 6.5833 | 7.3112 | 7 | 7.2333 | 7.7041 | 7.15 | 7 |  | 7.3833 | 6 | 6.9467 | 15.3333 | 5.5 | 7.2833 | 7.3167
| Median |  | 7.0872 | 6.63 | 7.3051 | 7 | 7.25 | 7.6743 | 7.15 | 7 |  | 7.35 | 6 | 6.95 | 15 | 5.5 | 7.3 | 7.4
| Std Dev. |  | 0.11291 | 0.11843 | 0.1083 | 0 | 0.24221 | 0.29611 | 0.16432 | 0 |  | 0.24833 | 0 | 0.087101 | 0.5164 | 0.54772 | 0.26394 | 0.13292
| Rel.Std.Dev. |  | 1.59% | 1.799% | 1.481% | 0% | 3.349% | 3.844% | 2.298% | 0% |  | 3.363% | 0% | 1.254% | 3.368% | 9.959% | 3.624% | 1.817%
| PDM3 |  | -0.9021% | -8.11% | 2.05% | -2.294% | 0.9624% | 7.533% | -0.2007% | -2.294% |  | 3.056% | -16.25% | -3.039% | 114% | -23.23% | 1.66% | 2.126%
| Z-Score (Absolute) |  | 0.17 | 1.88 | 0.53 | 0.5 | 0.28 | 1.56 | 0 | 0.5 | NA | 0.77 | 3.81 | 0.67 | 27.14 | 5.47 | 0.44 | 0.7
| Analytical results for Yb in OREAS 604 (Certified Value 0.56 ppm) | [23 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 09: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | [6 empty] | 1
| No | No | 01: 02; 02: 09; 09: 13; 11: 15; 12: 17; 13: 19; 14: 25 | [6 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 09: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*OES/MS; 14: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 01: 0.55; 02: 0.56738; 09: 0.39307; 11: 0.7; 12: 0.6; 13: 0.56; 14: 0.6 | [6 empty] | 1
| 1 | 2 | 01: 0.54; 02: 0.55055; 09: 0.40328; 11: 0.56; 12: 0.5; 13: 0.54; 14: 0.6 | [6 empty] | 0
| 1 | 3 | 01: 0.52; 02: 0.59415; 09: 0.41312; 11: 0.59; 12: 0.5; 13: 0.53; 14: 0.6 | [6 empty] | 16
| 1 | 4 | 01: 0.54; 02: 0.57345; 09: 0.40487; 11: 0.62; 12: 0.5; 13: 0.57; 14: 0.75 | [6 empty] | 0.5631
| 1 | 5 | 01: 0.54; 02: 0.58257; 09: 0.4081; 11: 0.57; 12: 0.6; 13: 0.55; 14: 0.65 | [6 empty] | 65
| 1 | 6 | 01: 0.55; 02: 0.55552; 09: 0.39122; 11: 0.6; 12: 0.5; 13: 0.53; 14: 0.55
| Mean |  | 01: 0.54; 02: 0.5706; 09: 0.40228; 11: 0.60667; 12: 0.53333; 13: 0.54667; 14: 0.625
| Median |  | 01: 0.54; 02: 0.57041; 09: 0.40407; 11: 0.595; 12: 0.5; 13: 0.545; 14: 0.6
| Std Dev. |  | 01: 0.010954; 02: 0.016408; 09: 0.0085589; 11: 0.050465; 12: 0.05164; 13: 0.01633; 14: 0.06892
| Rel.Std.Dev. |  | 01: 2.029%; 02: 2.876%; 09: 2.128%; 11: 8.318%; 12: 9.682%; 13: 2.987%; 14: 11.03%
| PDM3 |  | 01: -4.102%; 02: 1.332%; 09: -28.56%; 11: 7.737%; 12: -5.286%; 13: -2.918%; 14: 10.99%
| Z-Score (Absolute) |  | 01: 0.19; 02: 0.67; 09: 4.07; 11: 1.16; 12: 0.38; 13: 0; 14: 1.5
| Analytical results for Zn in OREAS 604 (Certified Value 0.255 wt.%) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 25 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.25938 | 0.26 | 0.262 | 0.252 | 0.265 | 0.252 | 0.26312 | 0.261 | 0.2679 | 0.2793 | 0.25246 | 0.2484 | 0.2578 | 0.2407 | 0.2397 | 0.2477 | 0.24897 | 0.2151 | 0.2466 | 0.247 | 0.254 |  | 1
| 1 | 2 | 0.25802 | 0.2621 | 0.258 | 0.261 | 0.261 | 0.254 | 0.26143 | 0.244 | 0.2647 | 0.2597 | 0.24999 | 0.2492 | 0.2571 | 0.2353 | 0.2416 | 0.2661 | 0.24831 | 0.2158 | 0.2452 | 0.244 | 0.257 |  | 0
| 1 | 3 | 0.25921 | 0.2609 | 0.26 | 0.256 | 0.262 | 0.247 | 0.25766 | 0.244 | 0.2674 | 0.2672 | 0.25337 | 0.2502 | 0.2615 | 0.23 | 0.2431 | 0.265 | 0.25495 | 0.2173 | 0.2485 | 0.244 | 0.252 |  | 16
| 1 | 4 | 0.25494 | 0.2597 | 0.258 | 0.252 | 0.259 | 0.251 | 0.26057 | 0.241 | 0.2732 | 0.266 | 0.25395 | 0.2472 | 0.2593 | 0.2273 | 0.2398 | 0.2639 | 0.26575 | 0.2084 | 0.2506 | 0.252 | 0.255 |  | 0.25485
| 1 | 5 | 0.25891 | 0.2625 | 0.253 | 0.277 | 0.255 | 0.251 | 0.26177 | 0.238 | 0.2678 | 0.2582 | 0.24627 | 0.2487 | 0.2576 | 0.2298 | 0.2445 | 0.2524 | 0.24517 | 0.218 | 0.2436 | 0.257 | 0.251 |  | 66
| 1 | 6 | 0.25745 | 0.268 | 0.254 | 0.255 | 0.266 | 0.256 | 0.26015 | 0.242 | 0.2631 | 0.2752 | 0.24913 | 0.2484 | 0.2609 | 0.2305 | 0.2413 | 0.2617 | 0.24762 | 0.2208 | 0.2452 | 0.256 | 0.25
| Mean |  | 0.25799 | 0.2622 | 0.2575 | 0.25883 | 0.26133 | 0.25183 | 0.26078 | 0.245 | 0.26735 | 0.2676 | 0.25086 | 0.24868 | 0.25903 | 0.23227 | 0.24167 | 0.25947 | 0.25179 | 0.2159 | 0.24662 | 0.25 | 0.25317
| Median |  | 0.25846 | 0.2615 | 0.258 | 0.2555 | 0.2615 | 0.2515 | 0.261 | 0.243 | 0.2676 | 0.2666 | 0.25123 | 0.24855 | 0.25855 | 0.23025 | 0.24145 | 0.2628 | 0.24864 | 0.21655 | 0.2459 | 0.2495 | 0.253
| Std Dev. |  | 0.0016641 | 0.0030502 | 0.0034496 | 0.0094956 | 0.0040332 | 0.0030605 | 0.0018466 | 0.0081486 | 0.0034588 | 0.0083443 | 0.0029402 | 0.00099281 | 0.0018414 | 0.0048854 | 0.0018747 | 0.0075849 | 0.0075666 | 0.004178 | 0.0025506 | 0.005831 | 0.0026394
| Rel.Std.Dev. |  | 0.645% | 1.163% | 1.34% | 3.669% | 1.543% | 1.215% | 0.7081% | 3.326% | 1.294% | 3.118% | 1.172% | 0.3992% | 0.7109% | 2.103% | 0.7758% | 2.923% | 3.005% | 1.935% | 1.034% | 2.332% | 1.043%
| PDM3 |  | 1.232% | 2.885% | 1.041% | 1.564% | 2.545% | -1.182% | 2.33% | -3.864% | 4.906% | 5.004% | -1.563% | -2.418% | 1.643% | -8.86% | -5.172% | 1.813% | -1.198% | -15.28% | -3.229% | -1.902% | -0.6592%
| Z-Score (Absolute) |  | 0.52 | 0.97 | 0.46 | 0.22 | 0.87 | 0.14 | 0.82 | 1.22 | 1.52 | 1.54 | 0.25 | 0.48 | 0.63 | 2.42 | 1.23 | 0.67 | 0.45 | 3.99 | 0.7 | 0.34 | 0
| Analytical results for Zr in OREAS 604 (Certified Value 104 ppm) | [23 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | [7 empty] | 0
0
| 1 | 1 | 109.48 | 100 | 110.7794 | 102 | 89.5 | 98.141 | 112.7 | 117.34 | 103 | 101 | 100 | 117.68 | 92 | 103 | 90 | [7 empty] | 1
| 1 | 2 | 111.2567 | 101.2 | 111.1049 | 107 | 83.5 | 99.0628 | 108.1 | 110.56 | 102 | 101 | 109 | 110.85 | 94 | 102 | 90 | [7 empty] | 0
| 1 | 3 | 108.875 | 100.7 | 111.5113 | 111 | 77.5 | 99.1108 | 109.7 | 103.69 | 102 | 105 | 108 | 109.44 | 93 | 101 | 93 | [7 empty] | 16
| 1 | 4 | 108.15 | 101.3 | 111.3811 | 115 | 90.3 | 101.0847 | 112.2 | 119.49 | 94 | 101 | 107 | 111.12 | 97 | 108 | 95 | [7 empty] | 104.4469
| 1 | 5 | 108.53 | 102.4 | 111.8951 | 108 | 79.8 | 98.3726 | 108.5 | 113.11 | 103 | 103 | 102 | 109.11 | 94 | 103 | 95 | [7 empty] | 67
| 1 | 6 | 108.965 | 104.6 | 111.9234 | 98 | 87 | 99.5017 | 112.3 | 123.4 | 99 | 104 | 100 | 109.32 | 90 | 109 | 95
| Mean |  | 109.2094 | 101.7 | 111.4325 | 106.8333 | 84.6 | 99.2123 | 110.5833 | 114.5983 | 100.5 | 102.5 | 104.3333 | 111.2533 | 93.3333 | 104.3333 | 93
| Median |  | 108.92 | 101.25 | 111.4462 | 107.5 | 85.25 | 99.0868 | 110.95 | 115.225 | 102 | 102 | 104.5 | 110.145 | 93.5 | 103 | 94
| Std Dev. |  | 1.0973 | 1.6248 | 0.44672 | 6.1128 | 5.233 | 1.0464 | 2.0653 | 7.0204 | 3.5071 | 1.7607 | 4.1312 | 3.2589 | 2.3381 | 3.3267 | 2.4495
| Rel.Std.Dev. |  | 1.005% | 1.598% | 0.4009% | 5.722% | 6.186% | 1.055% | 1.868% | 6.126% | 3.49% | 1.718% | 3.96% | 2.929% | 2.505% | 3.188% | 2.634%
| PDM3 |  | 4.56% | -2.63% | 6.688% | 2.285% | -19% | -5.012% | 5.875% | 9.719% | -3.779% | -1.864% | -0.1088% | 6.517% | -10.64% | -0.1088% | -10.96%
| Z-Score (Absolute) |  | 0.64 | 0.42 | 0.93 | 0.33 | 2.6 | 0.67 | 0.82 | 1.35 | 0.33 | 0.24 | 0 | 0.74 | 1.45 | 0 | 1.49
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 604 (Indicative Value 0.158 wt.%) | [17 empty] | Indicative
C | Round | Replicate | Lab | Lab | [14 empty] | 1
| No | No | 02 | 21 | [14 empty] | wt.%
|  |  | IRC | IRC | [14 empty] | 3
3
| 1 | 1 | 0.16 | 0.15831 | [14 empty] | 1
| 1 | 2 | 0.16 | 0.16318 | [14 empty] | 10
| 1 | 3 | 0.27 | 0.15502 | [14 empty] | 16
| 1 | 4 | 0.16 | 0.15667 | [14 empty] | 0.1578
| 1 | 5 | 0.16 | 0.15369 | [14 empty] | 16
| 1 | 6 | 0.15 | 0.15877
| Mean |  | 0.17667 | 0.1576
| Median |  | 0.16 | 0.15749
| Std Dev. |  | 0.045898 | 0.0033415
| Rel.Std.Dev. |  | 25.98% | 2.12%
| PDM3 |  | 11.95% | -0.1254%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 604 (Certified Value 4.85 wt.%) | [17 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 06 | 07 | 09 | 11 | 12 | 14 | 17 | 18 | 19 | 21 | 22 | 24 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 4.88 | 4.97 | 4.7 | 4.78 | 4.765 | 4.74 | 4.94 | 4.62 | 4.98 | 5.313 | 4.75 | 5.4913 | 4.58 | 5.1 | 4.42 | 5.2 | 1
| 1 | 2 | 4.83 | 4.93 | 4.69 | 4.57 | 4.757 | 4.74 | 5.04 | 4.61 | 5.04 | 5.379 | 4.71 | 5.5855 | 4.12 | 5.06 | 4.42 | 5.11 | 0
| 1 | 3 | 4.86 | 4.94 | 4.65 | 4.81 | 4.776 | 4.72 | 5.03 | 4.56 | 5.01 | 5.324 | 4.71 | 5.4658 |  | 4.93 | 4.75 | 5.16 | 16
| 1 | 4 | 4.85 | 4.93 | 4.63 | 4.76 | 4.791 | 4.87 | 4.97 | 4.61 | 4.92 | 5.258 | 4.76 | 5.4365 |  | 5.01 | 4.75 | 5.19 | 4.8539
| 1 | 5 | 4.83 | 4.94 | 4.68 | 4.77 | 4.8 | 4.74 | 4.96 | 4.46 | 5.09 | 5.302 | 4.76 | 5.3801 |  | 4.96 | 4.65 | 5.2 | 128
| 1 | 6 | 4.82 | 4.97 | 4.68 | 4.71 | 4.754 | 4.72 | 9.24 | 4.66 | 5.05 | 5.346 | 4.71 | 5.4038 |  | 4.96 | 4.74 | 5.2
| Mean |  | 4.845 | 4.9467 | 4.6717 | 4.7333 | 4.7738 | 4.755 | 5.6967 | 4.5867 | 5.015 | 5.3203 | 4.7333 | 5.4605 | 4.35 | 5.0033 | 4.6217 | 5.1767
| Median |  | 4.84 | 4.94 | 4.68 | 4.765 | 4.7705 | 4.74 | 5 | 4.61 | 5.025 | 5.3185 | 4.73 | 5.4511 | 4.35 | 4.985 | 4.695 | 5.195
| Std Dev. |  | 0.022583 | 0.018619 | 0.026394 | 0.08641 | 0.018649 | 0.057184 | 1.7363 | 0.069761 | 0.059582 | 0.040962 | 0.02582 | 0.07331 | 0.32527 | 0.065929 | 0.16068 | 0.036148
| Rel.Std.Dev. |  | 0.4661% | 0.3764% | 0.565% | 1.826% | 0.3906% | 1.203% | 30.48% | 1.521% | 1.188% | 0.7699% | 0.5455% | 1.343% | 7.477% | 1.318% | 3.477% | 0.6983%
| PDM3 |  | -0.1834% | 1.911% | -3.754% | -2.484% | -1.65% | -2.038% | 17.36% | -5.506% | 3.319% | 9.609% | -2.484% | 12.5% | -10.38% | 3.079% | -4.784% | 6.65%
| Z-Score (Absolute) |  | 0.13 | 0.51 | 0.51 | 0.16 | 0.13 | 0.2 | 0.66 | 0.73 | 0.76 | 1.88 | 0.28 | 2.4 | 0.84 | 0.71 | 0.69 | 1.35
Note on C17: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. Laboratory batches are rejected if Z-Score > 2.5.