## Abbreviations
Table 1. Abbreviations used for OREAS 45f
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
5A*MS | 5-acid (HF-HNO3-HClO4-HCl-H3BO3) digest with Mass Spectrometry: ICP-MS finish
5A*OES | 5-acid (HF-HNO3-HClO4-HCl-H3BO3) digest with Optical Emmision Spectrometry [aka: A(tomic)ES, ICP-OES] finish
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
at 105°C*Oven | moisture at 105°C with drying oven finish
at 105°C*TGA | moisture at 105°C with Thermal Gravimetric Analyser finish
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
FA*AAS | fire assay with AAS finish
FA*MS | fire assay with ICP-MS finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
LOI*Furnace | loss on ignition with muffle furnace finish
LOI*TGA | loss on ignition with Thermal Gravimetric Analyser finish
NiS*MS | fire assay (nickel sulphide) with Mass Spectrometry: ICP-MS finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Text Values:
N.A. | Not Applicable (Lab 34)
NA | Not Applicable (Lab 17)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 45f
Lab Slot | Laboratory Name and Location
AGAT Laboratories, Mississauga, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Perth, WA, Australia
Nagrom, Perth, WA, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
PT SGS Indo Assay Laboratories, Jakarta, Indonesia
SGS, Randfontein, Gauteng, South Africa
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Zarazma Mineral Studies Company, Tehran, Iran

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 45f
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 19.2844 | 1.7216 | 18.6301 | 19.9387 | 18.3491 | 20.2197
Pd, Palladium (ppb) | 56.562 | 3.9943 | 54.9506 | 58.1733 | 55.2698 | 57.8541
Pt, Platinum (ppb) | 38.079 | 2.4929 | 37.141 | 39.0169 | 36.2994 | 39.8585
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppb) | 18.0446 | 1.6253 | 17.3131 | 18.7761 | 17.0859 | 19.0033
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 20.4579 | 0.18739 | 20.3737 | 20.5422 | 20.3475 | 20.5683
BaO, Barium oxide (ppm) | 265.7438 | 45.3878 | 207.475 | 324.0126 | 239.911 | 291.5766
CaO, Calcium oxide (wt.%) | 0.1394 | 0.005249 | 0.13681 | 0.14198 | 0.13554 | 0.14325
Cr2O3, Chromium(III) oxide (ppm) | 706.8244 | 33.1091 | 695.4144 | 718.2345 | 661.1565 | 752.4923
Cu, Copper (ppm) | 364.1755 | 15.4472 | 348.5028 | 379.8483 | 353.1838 | 375.1673
Fe2O3, Iron(III) oxide (wt.%) | 22.3127 | 0.2941 | 22.1793 | 22.446 | 22.1856 | 22.4397
K2O, Potassium oxide (wt.%) | 0.28391 | 0.0059463 | 0.28148 | 0.28633 | 0.27989 | 0.28792
MgO, Magnesium oxide (wt.%) | 0.42251 | 0.0097042 | 0.41801 | 0.42702 | 0.41272 | 0.4323
MnO, Manganese oxide (wt.%) | 0.030848 | 0.0014744 | 0.030216 | 0.03148 | 0.028971 | 0.032725
Na2O, Sodium oxide (wt.%) | 0.091243 | 0.012809 | 0.085138 | 0.097349 | 0.084932 | 0.097555
Ni, Nickel (ppm) | 292.6923 | 23.0129 | 275.7027 | 309.682 | 267.0628 | 318.3219
P2O5, Phosphorus(V) oxide (wt.%) | 0.070773 | 0.0041559 | 0.067851 | 0.073695 | 0.068734 | 0.072812
S, Sulphur (wt.%) | 0.031501 | 0.0018146 | 0.029805 | 0.033197 | 0.02946 | 0.033542
SiO2, Silicon dioxide (wt.%) | 43.5888 | 0.25157 | 43.4698 | 43.7079 | 43.4078 | 43.7698
Sr, Strontium (ppm) | 31.3732 | 6.5306 | 27.0567 | 35.6897 | IND | IND
TiO2, Titanium dioxide (wt.%) | 1.9952 | 0.019756 | 1.9868 | 2.0036 | 1.9803 | 2.0101
V2O5, Vanadium(V) oxide (ppm) | 447.2114 | 55.9343 | 385.0533 | 509.3694 | 427.7903 | 466.6324
Zr, Zirconium (ppm) | 263.0773 | 17.604 | 245.0854 | 281.0691 | 255.0761 | 271.0784
Thermogravimetry
LOI1000, Loss on ignition @1000°C (wt.%) | 10.1255 | 0.21675 | 10.0348 | 10.2162 | 10.0431 | 10.2078
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 10.4102 | 0.21205 | 10.3156 | 10.5048 | 10.2546 | 10.5658
Ba, Barium (ppm) | 210.4902 | 9.5497 | 206.1351 | 214.8452 | 204.0629 | 216.9175
Be, Beryllium (ppm) | 1.0329 | 0.15252 | 0.86456 | 1.2013 | IND | IND
Bi, Bismuth (ppm) | 0.19993 | 0.039762 | 0.1739 | 0.22596 | IND | IND
Ca, Calcium (wt.%) | 0.10244 | 0.014545 | 0.09294 | 0.11194 | 0.093115 | 0.11177
Ce, Cerium (ppm) | 29.385 | 1.584 | 28.1995 | 30.5705 | 27.8736 | 30.8964
Co, Cobalt (ppm) | 49.0231 | 3.0225 | 47.0914 | 50.9549 | 46.7952 | 51.251
Cr, Chromium (ppm) | 475.5643 | 22.5253 | 462.4143 | 488.7143 | 459.1383 | 491.9903
Cs, Cesium (ppm) | 3.6605 | 0.2442 | 3.529 | 3.7921 | 3.4632 | 3.8578
Cu, Copper (ppm) | 361.5044 | 14.8961 | 353.806 | 369.2027 | 347.4855 | 375.5232
Dy, Dysprosium (ppm) | 2.8793 | 0.19344 | 2.7274 | 3.0311 | 2.6632 | 3.0953
Er, Erbium (ppm) | 1.6499 | 0.17449 | 1.5295 | 1.7703 | 1.5455 | 1.7543
Eu, Europium (ppm) | 0.62283 | 0.068244 | 0.57259 | 0.67308 | 0.57889 | 0.66678
Fe, Iron (wt.%) | 15.1571 | 0.35295 | 14.9973 | 15.3168 | 14.9506 | 15.3635
Ga, Gallium (ppm) | 26.1739 | 2.4558 | 24.5769 | 27.771 | 24.7194 | 27.6285
Gd, Gadolinium (ppm) | 2.5736 | 0.26072 | 2.363 | 2.7843 | 2.3341 | 2.8131
Hf, Hafnium (ppm) | 6.9683 | 0.78618 | 5.9461 | 7.9906 | 6.3929 | 7.5438
Ho, Holmium (ppm) | 0.56015 | 0.059697 | 0.51412 | 0.60618 | 0.51648 | 0.60381
K, Potassium (wt.%) | 0.22828 | 0.020101 | 0.21737 | 0.2392 | 0.21501 | 0.24156
La, Lanthanum (ppm) | 15.6891 | 1.6482 | 14.4571 | 16.9211 | 14.6987 | 16.6796
Li, Lithium (ppm) | 20.4394 | 1.3948 | 19.1406 | 21.7381 | 18.4849 | 22.3938
Lu, Lutetium (ppm) | 0.25858 | 0.035297 | 0.22984 | 0.28733 | 0.23231 | 0.28486
Mg, Magnesium (wt.%) | 0.24178 | 0.011199 | 0.23641 | 0.24715 | 0.23448 | 0.24909
Mn, Manganese (wt.%) | 0.023584 | 0.0018654 | 0.022347 | 0.02482 | 0.022293 | 0.024874
Na, Sodium (wt.%) | 0.069535 | 0.0073427 | 0.061384 | 0.077686 | IND | IND
Nb, Niobium (ppm) | 23.73 | 2.0392 | 22.6094 | 24.8505 | 22.6243 | 24.8356
Nd, Neodymium (ppm) | 12.9512 | 0.89582 | 12.1489 | 13.7534 | 12.264 | 13.6384
Ni, Nickel (ppm) | 268.4619 | 22.5353 | 255.892 | 281.0318 | 255.7647 | 281.159
P, Phosphorus (wt.%) | 0.029817 | 0.0043956 | 0.027362 | 0.032271 | 0.024621 | 0.035012
Pr, Praseodymium (ppm) | 3.4421 | 0.22138 | 3.2592 | 3.6251 | 3.1377 | 3.7466
Rb, Rubidium (ppm) | 31.5532 | 2.0245 | 30.2497 | 32.8566 | 30.4938 | 32.6126
Sb, Antimony (ppm) | 0.62922 | 0.14225 | 0.53247 | 0.72596 | IND | IND
Sc, Scandium (ppm) | 35.7957 | 2.0184 | 33.7427 | 37.8486 | 33.6824 | 37.9089
Si, Silicon (wt.%) | 19.7176 | 0.59827 | 19.4382 | 19.997 | 19.4416 | 19.9936
Sm, Samarium (ppm) | 2.6018 | 0.16042 | 2.4806 | 2.7229 | 2.4636 | 2.7399
Sn, Tin (ppm) | 3.369 | 0.67133 | 2.8114 | 3.9265 | IND | IND
Sr, Strontium (ppm) | 26.8555 | 4.2088 | 23.6319 | 30.0792 | 24.6337 | 29.0774
Ta, Tantalum (ppm) | 1.5551 | 0.14433 | 1.4621 | 1.6482 | 1.3984 | 1.7119
Tb, Terbium (ppm) | 0.4094 | 0.040486 | 0.38192 | 0.43688 | 0.37791 | 0.44089
Th, Thorium (ppm) | 9.9905 | 0.79371 | 9.5125 | 10.4685 | 9.5709 | 10.4101
Ti, Titanium (wt.%) | 1.1439 | 0.041459 | 1.1253 | 1.1625 | 1.1223 | 1.1655
Tm, Thulium (ppm) | 0.26176 | 0.03702 | 0.23117 | 0.29234 | 0.24308 | 0.28043
U, Uranium (ppm) | 2.3179 | 0.19079 | 2.21 | 2.4259 | 2.1833 | 2.4525
V, Vanadium (ppm) | 261.4327 | 7.3075 | 256.571 | 266.2944 | 254.0916 | 268.7737
Y, Yttrium (ppm) | 14.7068 | 1.4564 | 13.8588 | 15.5548 | 14.1305 | 15.283
Yb, Ytterbium (ppm) | 1.7942 | 0.16808 | 1.6946 | 1.8938 | 1.688 | 1.9004
Zn, Zinc (ppm) | 34.3387 | 5.2135 | 30.6092 | 38.0682 | 30.9077 | 37.7698
Zr, Zirconium (ppm) | 261.2643 | 20.2427 | 243.0288 | 279.4997 | 246.1489 | 276.3796
4-Acid Digestion
Al, Aluminium (wt.%) | 10.1599 | 0.45139 | 9.9991 | 10.3207 | 9.9287 | 10.391
As, Arsenic (ppm) | 9.6675 | 0.84242 | 9.4215 | 9.9136 | 8.9675 | 10.3676
Ba, Barium (ppm) | 206.434 | 9.1737 | 202.8831 | 209.9848 | 201.5712 | 211.2967
Be, Beryllium (ppm) | 1.2048 | 0.14744 | 1.159 | 1.2505 | 1.1294 | 1.2801
Bi, Bismuth (ppm) | 0.21265 | 0.016544 | 0.20441 | 0.22088 | 0.19834 | 0.22695
Ca, Calcium (wt.%) | 0.095873 | 0.0074057 | 0.093111 | 0.098634 | 0.093461 | 0.098285
Ce, Cerium (ppm) | 28.8195 | 2.3996 | 27.8555 | 29.7835 | 27.7724 | 29.8666
Co, Cobalt (ppm) | 44.5486 | 2.3169 | 43.7317 | 45.3655 | 43.5196 | 45.5776
Cr, Chromium (ppm) | 417.3037 | 34.5497 | 404.5009 | 430.1065 | 403.3762 | 431.2313
Cs, Cesium (ppm) | 3.6493 | 0.24763 | 3.5578 | 3.7408 | 3.5031 | 3.7955
Cu, Copper (ppm) | 362.5909 | 15.5916 | 357.072 | 368.1098 | 354.9644 | 370.2174
Dy, Dysprosium (ppm) | 2.2254 | 0.13162 | 2.1352 | 2.3155 | 2.1029 | 2.3478
Er, Erbium (ppm) | 1.3331 | 0.090264 | 1.2815 | 1.3848 | 1.2762 | 1.39
Eu, Europium (ppm) | 0.63351 | 0.074688 | 0.5854 | 0.68161 | 0.59928 | 0.66773
Fe, Iron (wt.%) | 14.6498 | 0.56085 | 14.4318 | 14.8678 | 14.4103 | 14.8893
Ga, Gallium (ppm) | 26.7491 | 1.2368 | 26.2966 | 27.2016 | 25.835 | 27.6632
Gd, Gadolinium (ppm) | 2.3082 | 0.26666 | 2.1261 | 2.4904 | 2.1918 | 2.4246
Hf, Hafnium (ppm) | 4.6375 | 0.43459 | 4.4758 | 4.7992 | 4.454 | 4.821
Ho, Holmium (ppm) | 0.44714 | 0.051802 | 0.41288 | 0.4814 | 0.42358 | 0.47071
In, Indium (ppm) | 0.10554 | 0.0099766 | 0.10266 | 0.10843 | 0.094299 | 0.11679
K, Potassium (wt.%) | 0.22436 | 0.0089434 | 0.22119 | 0.22752 | 0.21808 | 0.23064
La, Lanthanum (ppm) | 15.7192 | 1.4269 | 15.1986 | 16.2398 | 15.0915 | 16.347
Li, Lithium (ppm) | 20.4363 | 1.1809 | 19.9983 | 20.8743 | 19.3069 | 21.5657
Lu, Lutetium (ppm) | 0.18593 | 0.027837 | 0.17124 | 0.20062 | 0.17505 | 0.19681
Mg, Magnesium (wt.%) | 0.22932 | 0.01499 | 0.22373 | 0.23492 | 0.22317 | 0.23547
Mn, Manganese (wt.%) | 0.022116 | 0.0013743 | 0.021579 | 0.022654 | 0.021584 | 0.022649
Mo, Molybdenum (ppm) | 2.2702 | 0.21927 | 2.1866 | 2.3537 | 2.1709 | 2.3695
Na, Sodium (wt.%) | 0.062532 | 0.0041637 | 0.06032 | 0.064745 | 0.061267 | 0.063798
Nb, Niobium (ppm) | 23.1488 | 1.5356 | 22.5236 | 23.774 | 22.3008 | 23.9967
Nd, Neodymium (ppm) | 12.2541 | 0.93069 | 11.6972 | 12.811 | 11.8529 | 12.6552
Ni, Nickel (ppm) | 256.4411 | 15.6305 | 250.7936 | 262.0886 | 250.2618 | 262.6204
P, Phosphorus (wt.%) | 0.029632 | 0.0018525 | 0.028917 | 0.030346 | 0.028412 | 0.030852
Pb, Lead (ppm) | 14.7159 | 1.177 | 14.2787 | 15.153 | 14.1589 | 15.2728
Pr, Praseodymium (ppm) | 3.4311 | 0.27683 | 3.2301 | 3.632 | 3.2823 | 3.5798
Rb, Rubidium (ppm) | 31.2439 | 1.9412 | 30.4767 | 32.011 | 30.2002 | 32.2875
S, Sulphur (wt.%) | 0.029177 | 0.0017528 | 0.028552 | 0.029803 | 0.027973 | 0.030382
Sb, Antimony (ppm) | 0.63635 | 0.061612 | 0.60886 | 0.66385 | 0.59005 | 0.68266
Sc, Scandium (ppm) | 36.2753 | 2.2702 | 35.4365 | 37.1141 | 34.9918 | 37.5589
Se, Selenium (ppm) | 2.2628 | 0.3769 | 2.0713 | 2.4542 | 2.0617 | 2.4638
Sm, Samarium (ppm) | 2.4874 | 0.14574 | 2.4024 | 2.5724 | 2.3541 | 2.6208
Sn, Tin (ppm) | 2.8506 | 0.16515 | 2.7951 | 2.9061 | 2.6769 | 3.0243
Sr, Strontium (ppm) | 25.1016 | 1.2201 | 24.6883 | 25.5149 | 24.1957 | 26.0075
Ta, Tantalum (ppm) | 1.6552 | 0.24899 | 1.5563 | 1.754 | 1.5781 | 1.7322
Tb, Terbium (ppm) | 0.3716 | 0.02361 | 0.35889 | 0.38431 | 0.35091 | 0.39228
Th, Thorium (ppm) | 9.9911 | 0.89358 | 9.6524 | 10.3299 | 9.5986 | 10.3837
Ti, Titanium (wt.%) | 1.0765 | 0.039826 | 1.0621 | 1.0909 | 1.0487 | 1.1043
Tl, Thallium (ppm) | 0.2017 | 0.01654 | 0.19435 | 0.20906 | 0.19184 | 0.21157
Tm, Thulium (ppm) | 0.2024 | 0.029777 | 0.18331 | 0.22149 | 0.18261 | 0.22219
U, Uranium (ppm) | 2.0888 | 0.15049 | 2.034 | 2.1435 | 2.016 | 2.1615
V, Vanadium (ppm) | 252.5181 | 8.3831 | 249.5922 | 255.4439 | 247.469 | 257.5672
W, Tungsten (ppm) | 1.2688 | 0.1421 | 1.215 | 1.3226 | 1.2083 | 1.3293
Y, Yttrium (ppm) | 10.8951 | 1.2381 | 10.4271 | 11.363 | 10.3969 | 11.3933
Yb, Ytterbium (ppm) | 1.2476 | 0.13838 | 1.1768 | 1.3184 | 1.2005 | 1.2947
Zn, Zinc (ppm) | 35.2584 | 2.2603 | 34.3605 | 36.1563 | 33.801 | 36.7158
Zr, Zirconium (ppm) | 171.942 | 11.2396 | 167.7474 | 176.1365 | 165.8925 | 177.9914
Aqua Regia Digestion (sample weights 0.15-50g)
Al, Aluminium (wt.%) | 4.8054 | 0.6234 | 4.5314 | 5.0793 | 4.6839 | 4.9269
Ba, Barium (ppm) | 157.9904 | 6.538 | 155.387 | 160.5939 | 154.451 | 161.5298
Be, Beryllium (ppm) | 0.98214 | 0.090358 | 0.94856 | 1.0157 | 0.93792 | 1.0264
Bi, Bismuth (ppm) | 0.16801 | 0.013975 | 0.16119 | 0.17482 | 0.15833 | 0.17768
Ca, Calcium (wt.%) | 0.075053 | 0.0064004 | 0.072824 | 0.077283 | 0.073566 | 0.076541
Ce, Cerium (ppm) | 22.3091 | 2.2539 | 21.3999 | 23.2183 | 21.5682 | 23.05
Co, Cobalt (ppm) | 39.2404 | 3.4399 | 37.9395 | 40.5414 | 38.1564 | 40.3244
Cr, Chromium (ppm) | 341.1928 | 25.3831 | 331.6056 | 350.78 | 333.3346 | 349.051
Cs, Cesium (ppm) | 1.876 | 0.2659 | 1.7471 | 2.005 | 1.7798 | 1.9723
Cu, Copper (ppm) | 335.8292 | 16.2292 | 329.9441 | 341.7142 | 326.7069 | 344.9515
Dy, Dysprosium (ppm) | 1.4915 | 0.11343 | 1.4094 | 1.5735 | 1.4299 | 1.553
Er, Erbium (ppm) | 0.78286 | 0.063233 | 0.7369 | 0.82881 | 0.74852 | 0.8172
Eu, Europium (ppm) | 0.49195 | 0.060638 | 0.44618 | 0.53771 | 0.46645 | 0.51744
Fe, Iron (wt.%) | 13.6892 | 0.55959 | 13.4773 | 13.9012 | 13.4044 | 13.974
Ga, Gallium (ppm) | 20.3281 | 1.9143 | 19.5134 | 21.1427 | 19.5587 | 21.0974
Gd, Gadolinium (ppm) | 1.7022 | 0.15918 | 1.5838 | 1.8206 | 1.6273 | 1.7771
Ge, Germanium (ppm) | 0.11746 | 0.017182 | 0.099988 | 0.13492 | 0.10777 | 0.12714
Hf, Hafnium (ppm) | 0.92755 | 0.16557 | 0.84672 | 1.0084 | 0.87746 | 0.97765
Hg, Mercury (ppm) | 0.030726 | 0.0050053 | 0.028237 | 0.033216 | 0.024937 | 0.036516
Ho, Holmium (ppm) | 0.27807 | 0.022204 | 0.26059 | 0.29555 | 0.26655 | 0.28959
In, Indium (ppm) | 0.087199 | 0.0064147 | 0.084578 | 0.089821 | 0.083104 | 0.091295
K, Potassium (wt.%) | 0.082108 | 0.014102 | 0.076213 | 0.088002 | 0.079027 | 0.085189
La, Lanthanum (ppm) | 10.7474 | 0.74942 | 10.4314 | 11.0634 | 10.4412 | 11.0536
Lu, Lutetium (ppm) | 0.097035 | 0.0068265 | 0.093476 | 0.10059 | 0.0894 | 0.10467
Mg, Magnesium (wt.%) | 0.15205 | 0.020926 | 0.14406 | 0.16003 | 0.14654 | 0.15756
Mn, Manganese (wt.%) | 0.014642 | 0.0016359 | 0.013988 | 0.015296 | 0.014162 | 0.015122
Mo, Molybdenum (ppm) | 1.1928 | 0.22622 | 1.0976 | 1.288 | 1.1297 | 1.2559
Na, Sodium (wt.%) | 0.032282 | 0.0037512 | 0.030207 | 0.034358 | 0.031197 | 0.033368
Nd, Neodymium (ppm) | 10.0694 | 1.0122 | 9.3436 | 10.7953 | 9.6791 | 10.4598
Ni, Nickel (ppm) | 191.8644 | 13.6796 | 186.0897 | 197.639 | 185.9561 | 197.7727
P, Phosphorus (wt.%) | 0.022145 | 0.0014259 | 0.021672 | 0.022619 | 0.021006 | 0.023285
Pb, Lead (ppm) | 12.3904 | 0.59958 | 12.187 | 12.5939 | 12.0375 | 12.7434
Pd, Palladium (ppb) | 39.9213 | 2.2634 | 39.2328 | 40.6097 | 35.4804 | 44.3621
Pr, Praseodymium (ppm) | 2.6333 | 0.3665 | 2.3416 | 2.9249 | 2.538 | 2.7286
Pt, Platinum (ppb) | 36.653 | 2.012 | 35.4604 | 37.8455 | 33.2619 | 40.0441
Rb, Rubidium (ppm) | 14.3727 | 1.0771 | 13.8526 | 14.8928 | 13.8256 | 14.9198
S, Sulphur (wt.%) | 0.026561 | 0.0040637 | 0.024944 | 0.028179 | 0.025829 | 0.027293
Sc, Scandium (ppm) | 31.4172 | 2.2581 | 30.5629 | 32.2715 | 30.329 | 32.5054
Sm, Samarium (ppm) | 1.9094 | 0.24512 | 1.717 | 2.1018 | 1.8096 | 2.0092
Sn, Tin (ppm) | 1.9655 | 0.20629 | 1.8696 | 2.0613 | 1.8802 | 2.0508
Sr, Strontium (ppm) | 13.1565 | 1.1115 | 12.688 | 13.6251 | 12.658 | 13.655
Tb, Terbium (ppm) | 0.24841 | 0.014473 | 0.23908 | 0.25774 | 0.23361 | 0.2632
Th, Thorium (ppm) | 7.6691 | 0.6611 | 7.3989 | 7.9393 | 7.4721 | 7.8661
Ti, Titanium (wt.%) | 0.097419 | 0.019964 | 0.088182 | 0.10666 | 0.092616 | 0.10222
Tl, Thallium (ppm) | 0.11538 | 0.01209 | 0.1095 | 0.12125 | 0.11056 | 0.12019
Tm, Thulium (ppm) | 0.10731 | 0.0081412 | 0.098913 | 0.11571 | IND | IND
U, Uranium (ppm) | 1.0877 | 0.090751 | 1.051 | 1.1244 | 1.0559 | 1.1195
V, Vanadium (ppm) | 217.0678 | 10.5094 | 213.2223 | 220.9132 | 213.0859 | 221.0496
Y, Yttrium (ppm) | 6.7424 | 0.33559 | 6.6129 | 6.8718 | 6.4509 | 7.0338
Yb, Ytterbium (ppm) | 0.68935 | 0.045207 | 0.66616 | 0.71254 | 0.66516 | 0.71355
Zn, Zinc (ppm) | 22.2014 | 3.6624 | 20.661 | 23.7417 | 21.3471 | 23.0556
Zr, Zirconium (ppm) | 30.0433 | 4.2754 | 28.1411 | 31.9455 | 28.1945 | 31.8921
Infrared Combustion
C, Carbon (wt.%) | 0.31845 | 0.020695 | 0.31033 | 0.32658 | 0.30843 | 0.32848

## Indicative Values
Table 3. Indicative Values for OREAS 45f
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
NiS Fire Assay
Au | ppb | 17.6667 | Pd | ppb | 55.9819 | Rh | ppb | 3.3187
Ir | ppb | 1.4462 | Pt | ppb | 36.588 | Ru | ppb | 3.3065
Borate Fusion XRF
As | ppm | 9.4017 | F | ppm | 1478.0416 | Zn | ppm | 26.3593
Cl | ppm | 193.0409 | Rb | ppm | 36.8733
Co | ppm | 43.6747 | Sn | ppm | < 10
Thermogravimetry
H2O- | wt.% | 3.9513
Borate / Peroxide Fusion ICP
Ag | ppm | < 1 | Ir | ppb | < 15 | S | wt.% | 0.034351
As | ppm | 9.4889 | Mo | ppm | 2.1527 | Se | ppm | < 20
B | ppm | < 20 | Pb | ppm | 14.8667 | Te | ppm | < 1
Cd | ppm | < 0.2 | Pt | ppb | < 50 | Tl | ppm | < 0.5
Ge | ppm | 2.2454 | Re | ppm | < 0.01 | W | ppm | 1.374
I | ppm | 15.8333 | Rh | ppb | < 30
In | ppm | < 0.2 | Ru | ppb | < 30
4-Acid Digestion
Ag | ppm | 0.088056 | I | ppm | 14.3333 | Rh | ppb | < 5
B | ppm | 9.4742 | Ir | ppb | < 300 | Ru | ppb | < 10
Cd | ppm | 0.02433 | Pd | ppb | 59.3333 | Te | ppm | 0.080383
Ge | ppm | 0.27889 | Pt | ppb | 41.5
Hg | ppm | 0.033167 | Re | ppm | < 0.002
Aqua Regia Digestion (sample weights 0.15-50g)
Ag | ppm | 0.05665 | Li | ppm | 9.1922 | Se | ppm | 0.57696
As | ppm | 2.7 | Nb | ppm | 0.16544 | Ta | ppm | < 0.01
B | ppm | < 10 | Re | ppm | < 0.001 | Te | ppm | 0.03218
Cd | ppm | 0.011154 | Rh | ppb | < 10 | W | ppm | 0.0063381
I | ppm | 15.8333 | Ru | ppb | < 10
Ir | ppb | 7.2204 | Sb | ppm | 0.24367
Infrared Combustion
S | wt.% | 0.024832

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 45f
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 19.2844 | 1.7216 | 15.8412 | 22.7276 | 14.1197 | 24.4491 | 8.927% | 17.85% | 26.78% | 18.3202 | 20.2486
Pd, ppb | 56.562 | 3.9943 | 48.5733 | 64.5507 | 44.5789 | 68.545 | 7.062% | 14.12% | 21.19% | 53.7339 | 59.3901
Pt, ppb | 38.079 | 2.4929 | 33.0931 | 43.0648 | 30.6002 | 45.5577 | 6.547% | 13.09% | 19.64% | 36.175 | 39.9829
Aqua Regia Digestion (sample weights 10-50g)
Au, ppb | 18.0446 | 1.6253 | 14.794 | 21.2952 | 13.1687 | 22.9205 | 9.007% | 18.01% | 27.02% | 17.1424 | 18.9468
Borate Fusion XRF
Al2O3, wt.% | 20.4579 | 0.18739 | 20.0832 | 20.8327 | 19.8958 | 21.0201 | 0.916% | 1.832% | 2.748% | 19.435 | 21.4808
BaO, ppm | 265.7438 | 45.3878 | 174.9682 | 356.5194 | 129.5804 | 401.9072 | 17.08% | 34.16% | 51.24% | 252.4566 | 279.031
CaO, wt.% | 0.1394 | 0.005249 | 0.1289 | 0.14989 | 0.12365 | 0.15514 | 3.766% | 7.531% | 11.3% | 0.13243 | 0.14636
Cr2O3, ppm | 706.8244 | 33.1091 | 640.6062 | 773.0427 | 607.4971 | 806.1518 | 4.684% | 9.368% | 14.05% | 671.4832 | 742.1657
Cu, ppm | 364.1755 | 15.4472 | 333.2812 | 395.0699 | 317.834 | 410.5171 | 4.242% | 8.483% | 12.73% | 345.9668 | 382.3843
Fe2O3, wt.% | 22.3127 | 0.2941 | 21.7245 | 22.9009 | 21.4304 | 23.195 | 1.318% | 2.636% | 3.954% | 21.197 | 23.4283
K2O, wt.% | 0.28391 | 0.0059463 | 0.27202 | 0.2958 | 0.26607 | 0.30175 | 2.094% | 4.189% | 6.283% | 0.26971 | 0.2981
MgO, wt.% | 0.42251 | 0.0097042 | 0.40311 | 0.44192 | 0.3934 | 0.45163 | 2.297% | 4.594% | 6.89% | 0.40139 | 0.44364
MnO, wt.% | 0.030848 | 0.0014744 | 0.027899 | 0.033797 | 0.026425 | 0.035271 | 4.78% | 9.559% | 14.34% | 0.029305 | 0.03239
Na2O, wt.% | 0.091243 | 0.012809 | 0.065626 | 0.11686 | 0.052818 | 0.12967 | 14.04% | 28.08% | 42.11% | 0.086681 | 0.095806
Ni, ppm | 292.6923 | 23.0129 | 246.6665 | 338.7181 | 223.6536 | 361.7311 | 7.862% | 15.72% | 23.59% | 278.0577 | 307.327
P2O5, wt.% | 0.070773 | 0.0041559 | 0.062461 | 0.079085 | 0.058305 | 0.083241 | 5.872% | 11.74% | 17.62% | 0.067234 | 0.074312
S, wt.% | 0.031501 | 0.0018146 | 0.027872 | 0.03513 | 0.026057 | 0.036945 | 5.76% | 11.52% | 17.28% | 0.029926 | 0.033076
SiO2, wt.% | 43.5888 | 0.25157 | 43.0857 | 44.0919 | 42.8341 | 44.3435 | 0.5771% | 1.154% | 1.731% | 41.4094 | 45.7683
Sr, ppm | 31.3732 | 6.5306 | 18.312 | 44.4344 | 11.7814 | 50.9651 | 20.82% | 41.63% | 62.45% | 29.8045 | 32.9419
TiO2, wt.% | 1.9952 | 0.019756 | 1.9557 | 2.0347 | 1.9359 | 2.0545 | 0.9902% | 1.98% | 2.971% | 1.8954 | 2.095
V2O5, ppm | 447.2114 | 55.9343 | 335.3428 | 559.0799 | 279.4085 | 615.0142 | 12.51% | 25.01% | 37.52% | 424.8508 | 469.5719
Zr, ppm | 263.0773 | 17.604 | 227.8692 | 298.2853 | 210.2651 | 315.8894 | 6.692% | 13.38% | 20.07% | 249.9234 | 276.2311
Thermogravimetry
LOI1000, wt.% | 10.1255 | 0.21675 | 9.692 | 10.559 | 9.4752 | 10.7757 | 2.141% | 4.281% | 6.422% | 9.6192 | 10.6317
Borate / Peroxide Fusion ICP
Al, wt.% | 10.4102 | 0.21205 | 9.9861 | 10.8343 | 9.7741 | 11.0464 | 2.037% | 4.074% | 6.111% | 9.8897 | 10.9307
Ba, ppm | 210.4902 | 9.5497 | 191.3907 | 229.5896 | 181.841 | 239.1393 | 4.537% | 9.074% | 13.61% | 199.9656 | 221.0147
Be, ppm | 1.0329 | 0.15252 | 0.72788 | 1.338 | 0.57536 | 1.4905 | 14.77% | 29.53% | 44.3% | 0.98127 | 1.0846
Bi, ppm | 0.19993 | 0.039762 | 0.1204 | 0.27945 | 0.080641 | 0.31922 | 19.89% | 39.78% | 59.66% | 0.18993 | 0.20992
Ca, wt.% | 0.10244 | 0.014545 | 0.073352 | 0.13153 | 0.058808 | 0.14608 | 14.2% | 28.4% | 42.59% | 0.09732 | 0.10756
Ce, ppm | 29.385 | 1.584 | 26.2169 | 32.5531 | 24.6329 | 34.1371 | 5.391% | 10.78% | 16.17% | 27.9157 | 30.8543
Co, ppm | 49.0231 | 3.0225 | 42.978 | 55.0682 | 39.9555 | 58.0908 | 6.166% | 12.33% | 18.5% | 46.572 | 51.4743
Cr, ppm | 475.5643 | 22.5253 | 430.5136 | 520.615 | 407.9883 | 543.1404 | 4.737% | 9.473% | 14.21% | 451.7861 | 499.3425
Cs, ppm | 3.6605 | 0.2442 | 3.1721 | 4.1489 | 2.9279 | 4.3931 | 6.671% | 13.34% | 20.01% | 3.4775 | 3.8435
Cu, ppm | 361.5044 | 14.8961 | 331.7121 | 391.2966 | 316.816 | 406.1928 | 4.121% | 8.241% | 12.36% | 343.4292 | 379.5796
Dy, ppm | 2.8793 | 0.19344 | 2.4924 | 3.2662 | 2.299 | 3.4596 | 6.718% | 13.44% | 20.16% | 2.7353 | 3.0232
Er, ppm | 1.6499 | 0.17449 | 1.3009 | 1.9989 | 1.1264 | 2.1734 | 10.58% | 21.15% | 31.73% | 1.5674 | 1.7324
Eu, ppm | 0.62283 | 0.068244 | 0.48635 | 0.75932 | 0.4181 | 0.82757 | 10.96% | 21.91% | 32.87% | 0.59169 | 0.65397
Fe, wt.% | 15.1571 | 0.35295 | 14.4512 | 15.863 | 14.0982 | 16.2159 | 2.329% | 4.657% | 6.986% | 14.3992 | 15.9149
Ga, ppm | 26.1739 | 2.4558 | 21.2623 | 31.0856 | 18.8065 | 33.5414 | 9.383% | 18.77% | 28.15% | 24.8652 | 27.4826
Gd, ppm | 2.5736 | 0.26072 | 2.0522 | 3.095 | 1.7915 | 3.3558 | 10.13% | 20.26% | 30.39% | 2.4449 | 2.7023
Hf, ppm | 6.9683 | 0.78618 | 5.396 | 8.5407 | 4.6098 | 9.3269 | 11.28% | 22.56% | 33.85% | 6.6199 | 7.3167
Ho, ppm | 0.56015 | 0.059697 | 0.44075 | 0.67954 | 0.38106 | 0.73924 | 10.66% | 21.31% | 31.97% | 0.53214 | 0.58816
K, wt.% | 0.22828 | 0.020101 | 0.18808 | 0.26849 | 0.16798 | 0.28859 | 8.805% | 17.61% | 26.42% | 0.21687 | 0.2397
La, ppm | 15.6891 | 1.6482 | 12.3928 | 18.9855 | 10.7446 | 20.6336 | 10.51% | 21.01% | 31.52% | 14.9047 | 16.4736
Li, ppm | 20.4394 | 1.3948 | 17.6498 | 23.2289 | 16.255 | 24.6237 | 6.824% | 13.65% | 20.47% | 19.4174 | 21.4613
Lu, ppm | 0.25858 | 0.035297 | 0.18799 | 0.32918 | 0.15269 | 0.36448 | 13.65% | 27.3% | 40.95% | 0.24565 | 0.27151
Mg, wt.% | 0.24178 | 0.011199 | 0.21938 | 0.26418 | 0.20818 | 0.27538 | 4.632% | 9.264% | 13.9% | 0.22969 | 0.25387
Mn, wt.% | 0.023584 | 0.0018654 | 0.019853 | 0.027314 | 0.017987 | 0.02918 | 7.91% | 15.82% | 23.73% | 0.022404 | 0.024763
Na, wt.% | 0.069535 | 0.0073427 | 0.05485 | 0.084221 | 0.047507 | 0.091563 | 10.56% | 21.12% | 31.68% | 0.066058 | 0.073012
Nb, ppm | 23.73 | 2.0392 | 19.6515 | 27.8084 | 17.6123 | 29.8476 | 8.593% | 17.19% | 25.78% | 22.5435 | 24.9165
Nd, ppm | 12.9512 | 0.89582 | 11.1595 | 14.7428 | 10.2637 | 15.6387 | 6.917% | 13.83% | 20.75% | 12.3036 | 13.5988
Ni, ppm | 268.4619 | 22.5353 | 223.3913 | 313.5325 | 200.856 | 336.0678 | 8.394% | 16.79% | 25.18% | 255.0388 | 281.885
P, wt.% | 0.029817 | 0.0043956 | 0.021026 | 0.038608 | 0.01663 | 0.043003 | 14.74% | 29.48% | 44.23% | 0.028326 | 0.031308
Pr, ppm | 3.4421 | 0.22138 | 2.9994 | 3.8849 | 2.778 | 4.1063 | 6.431% | 12.86% | 19.29% | 3.27 | 3.6143
Rb, ppm | 31.5532 | 2.0245 | 27.5042 | 35.6021 | 25.4797 | 37.6266 | 6.416% | 12.83% | 19.25% | 29.9755 | 33.1308
Sb, ppm | 0.62922 | 0.14225 | 0.34472 | 0.91372 | 0.20247 | 1.056 | 22.61% | 45.21% | 67.82% | 0.59776 | 0.66068
Sc, ppm | 35.7957 | 2.0184 | 31.7588 | 39.8325 | 29.7404 | 41.8509 | 5.639% | 11.28% | 16.92% | 34.0059 | 37.5855
Si, wt.% | 19.7176 | 0.59827 | 18.5211 | 20.9141 | 17.9228 | 21.5124 | 3.034% | 6.068% | 9.103% | 18.7317 | 20.7035
Sm, ppm | 2.6018 | 0.16042 | 2.2809 | 2.9226 | 2.1205 | 3.083 | 6.166% | 12.33% | 18.5% | 2.4717 | 2.7319
Sn, ppm | 3.369 | 0.67133 | 2.0263 | 4.7116 | 1.355 | 5.383 | 19.93% | 39.85% | 59.78% | 3.2005 | 3.5374
Sr, ppm | 26.8555 | 4.2088 | 18.4378 | 35.2732 | 14.229 | 39.4821 | 15.67% | 31.34% | 47.02% | 25.5127 | 28.1983
Ta, ppm | 1.5551 | 0.14433 | 1.2665 | 1.8438 | 1.1221 | 1.9881 | 9.281% | 18.56% | 27.84% | 1.4774 | 1.6329
Tb, ppm | 0.4094 | 0.040486 | 0.32843 | 0.49037 | 0.28794 | 0.53086 | 9.889% | 19.78% | 29.67% | 0.38893 | 0.42987
Th, ppm | 9.9905 | 0.79371 | 8.4031 | 11.5779 | 7.6093 | 12.3716 | 7.945% | 15.89% | 23.83% | 9.491 | 10.49
Ti, wt.% | 1.1439 | 0.041459 | 1.061 | 1.2268 | 1.0195 | 1.2682 | 3.624% | 7.249% | 10.87% | 1.0867 | 1.2011
Tm, ppm | 0.26176 | 0.03702 | 0.18772 | 0.3358 | 0.1507 | 0.37282 | 14.14% | 28.29% | 42.43% | 0.24867 | 0.27484
U, ppm | 2.3179 | 0.19079 | 1.9364 | 2.6995 | 1.7456 | 2.8903 | 8.231% | 16.46% | 24.69% | 2.202 | 2.4338
V, ppm | 261.4327 | 7.3075 | 246.8177 | 276.0476 | 239.5102 | 283.3551 | 2.795% | 5.59% | 8.385% | 248.361 | 274.5043
Y, ppm | 14.7068 | 1.4564 | 11.7941 | 17.6195 | 10.3377 | 19.0758 | 9.903% | 19.81% | 29.71% | 13.9714 | 15.4421
Yb, ppm | 1.7942 | 0.16808 | 1.4581 | 2.1304 | 1.29 | 2.2985 | 9.368% | 18.74% | 28.1% | 1.7045 | 1.8839
Zn, ppm | 34.3387 | 5.2135 | 23.9118 | 44.7657 | 18.6983 | 49.9791 | 15.18% | 30.36% | 45.55% | 32.6218 | 36.0557
Zr, ppm | 261.2643 | 20.2427 | 220.7788 | 301.7498 | 200.5361 | 321.9925 | 7.748% | 15.5% | 23.24% | 248.2011 | 274.3275
4-Acid Digestion
Al, wt.% | 10.1599 | 0.45139 | 9.2571 | 11.0627 | 8.8057 | 11.5141 | 4.443% | 8.886% | 13.33% | 9.6519 | 10.6679
As, ppm | 9.6675 | 0.84242 | 7.9827 | 11.3524 | 7.1403 | 12.1948 | 8.714% | 17.43% | 26.14% | 9.1842 | 10.1509
Ba, ppm | 206.434 | 9.1737 | 188.0866 | 224.7814 | 178.9129 | 233.9551 | 4.444% | 8.888% | 13.33% | 196.1123 | 216.7557
Be, ppm | 1.2048 | 0.14744 | 0.90989 | 1.4996 | 0.76245 | 1.6471 | 12.24% | 24.48% | 36.71% | 1.1445 | 1.265
Bi, ppm | 0.21265 | 0.016544 | 0.17956 | 0.24573 | 0.16301 | 0.26228 | 7.78% | 15.56% | 23.34% | 0.20201 | 0.22328
Ca, wt.% | 0.095873 | 0.0074057 | 0.081061 | 0.11068 | 0.073656 | 0.11809 | 7.724% | 15.45% | 23.17% | 0.091079 | 0.10067
Ce, ppm | 28.8195 | 2.3996 | 24.0203 | 33.6187 | 21.6207 | 36.0183 | 8.326% | 16.65% | 24.98% | 27.3785 | 30.2605
Co, ppm | 44.5486 | 2.3169 | 39.9148 | 49.1824 | 37.5979 | 51.4993 | 5.201% | 10.4% | 15.6% | 42.3212 | 46.776
Cr, ppm | 417.3037 | 34.5497 | 348.2044 | 486.4031 | 313.6547 | 520.9527 | 8.279% | 16.56% | 24.84% | 396.4385 | 438.1689
Cs, ppm | 3.6493 | 0.24763 | 3.154 | 4.1446 | 2.9064 | 4.3922 | 6.786% | 13.57% | 20.36% | 3.4668 | 3.8318
Cu, ppm | 362.5909 | 15.5916 | 331.4077 | 393.7741 | 315.8161 | 409.3657 | 4.3% | 8.6% | 12.9% | 344.4614 | 380.7205
Dy, ppm | 2.2254 | 0.13162 | 1.9621 | 2.4886 | 1.8305 | 2.6202 | 5.915% | 11.83% | 17.74% | 2.1141 | 2.3366
Er, ppm | 1.3331 | 0.090264 | 1.1526 | 1.5137 | 1.0623 | 1.6039 | 6.771% | 13.54% | 20.31% | 1.2665 | 1.3998
Eu, ppm | 0.63351 | 0.074688 | 0.48413 | 0.78288 | 0.40944 | 0.85757 | 11.79% | 23.58% | 35.37% | 0.60183 | 0.66518
Fe, wt.% | 14.6498 | 0.56085 | 13.5281 | 15.7715 | 12.9672 | 16.3323 | 3.828% | 7.657% | 11.49% | 13.9173 | 15.3823
Ga, ppm | 26.7491 | 1.2368 | 24.2756 | 29.2226 | 23.0388 | 30.4594 | 4.624% | 9.247% | 13.87% | 25.4116 | 28.0866
Gd, ppm | 2.3082 | 0.26666 | 1.7749 | 2.8416 | 1.5083 | 3.1082 | 11.55% | 23.11% | 34.66% | 2.1928 | 2.4236
Hf, ppm | 4.6375 | 0.43459 | 3.7683 | 5.5067 | 3.3338 | 5.9413 | 9.371% | 18.74% | 28.11% | 4.4056 | 4.8694
Ho, ppm | 0.44714 | 0.051802 | 0.34354 | 0.55074 | 0.29174 | 0.60255 | 11.59% | 23.17% | 34.76% | 0.42478 | 0.4695
In, ppm | 0.10554 | 0.0099766 | 0.08559 | 0.1255 | 0.075614 | 0.13547 | 9.453% | 18.91% | 28.36% | 0.10027 | 0.11082
K, wt.% | 0.22436 | 0.0089434 | 0.20647 | 0.24225 | 0.19753 | 0.25119 | 3.986% | 7.972% | 11.96% | 0.21314 | 0.23558
La, ppm | 15.7192 | 1.4269 | 12.8654 | 18.573 | 11.4385 | 19.9999 | 9.077% | 18.15% | 27.23% | 14.9333 | 16.5052
Li, ppm | 20.4363 | 1.1809 | 18.0745 | 22.798 | 16.8937 | 23.9789 | 5.778% | 11.56% | 17.33% | 19.4145 | 21.4581
Lu, ppm | 0.18593 | 0.027837 | 0.13026 | 0.24161 | 0.10242 | 0.26944 | 14.97% | 29.94% | 44.92% | 0.17663 | 0.19523
Mg, wt.% | 0.22932 | 0.01499 | 0.19934 | 0.2593 | 0.18435 | 0.27429 | 6.537% | 13.07% | 19.61% | 0.21785 | 0.24079
Mn, wt.% | 0.022116 | 0.0013743 | 0.019368 | 0.024865 | 0.017993 | 0.026239 | 6.214% | 12.43% | 18.64% | 0.021011 | 0.023222
Mo, ppm | 2.2702 | 0.21927 | 1.8316 | 2.7087 | 1.6124 | 2.928 | 9.659% | 19.32% | 28.98% | 2.1567 | 2.3837
Na, wt.% | 0.062532 | 0.0041637 | 0.054205 | 0.07086 | 0.050041 | 0.075023 | 6.658% | 13.32% | 19.98% | 0.059406 | 0.065659
Nb, ppm | 23.1488 | 1.5356 | 20.0776 | 26.2199 | 18.5421 | 27.7554 | 6.633% | 13.27% | 19.9% | 21.9913 | 24.3062
Nd, ppm | 12.2541 | 0.93069 | 10.3927 | 14.1155 | 9.462 | 15.0462 | 7.595% | 15.19% | 22.78% | 11.6414 | 12.8668
Ni, ppm | 256.4411 | 15.6305 | 225.1801 | 287.7021 | 209.5495 | 303.3326 | 6.095% | 12.19% | 18.29% | 243.619 | 269.2631
P, wt.% | 0.029632 | 0.0018525 | 0.025927 | 0.033337 | 0.024074 | 0.035189 | 6.252% | 12.5% | 18.76% | 0.02815 | 0.031113
Pb, ppm | 14.7159 | 1.177 | 12.3618 | 17.0699 | 11.1848 | 18.2469 | 7.998% | 16% | 23.99% | 13.9801 | 15.4517
Pr, ppm | 3.4311 | 0.27683 | 2.8774 | 3.9847 | 2.6006 | 4.2616 | 8.068% | 16.14% | 24.21% | 3.2595 | 3.6026
Rb, ppm | 31.2439 | 1.9412 | 27.3615 | 35.1262 | 25.4204 | 37.0674 | 6.213% | 12.43% | 18.64% | 29.6817 | 32.8061
S, wt.% | 0.029177 | 0.0017528 | 0.025672 | 0.032683 | 0.023919 | 0.034436 | 6.007% | 12.01% | 18.02% | 0.027718 | 0.030636
Sb, ppm | 0.63635 | 0.061612 | 0.51313 | 0.75958 | 0.45152 | 0.82119 | 9.682% | 19.36% | 29.05% | 0.60454 | 0.66817
Sc, ppm | 36.2753 | 2.2702 | 31.7349 | 40.8157 | 29.4647 | 43.0859 | 6.258% | 12.52% | 18.77% | 34.4616 | 38.0891
Se, ppm | 2.2628 | 0.3769 | 1.509 | 3.0166 | 1.1321 | 3.3935 | 16.66% | 33.31% | 49.97% | 2.1496 | 2.3759
Sm, ppm | 2.4874 | 0.14574 | 2.1959 | 2.7789 | 2.0502 | 2.9247 | 5.859% | 11.72% | 17.58% | 2.3631 | 2.6118
Sn, ppm | 2.8506 | 0.16515 | 2.5203 | 3.1809 | 2.3551 | 3.3461 | 5.794% | 11.59% | 17.38% | 2.7081 | 2.9931
Sr, ppm | 25.1016 | 1.2201 | 22.6615 | 27.5417 | 21.4415 | 28.7618 | 4.86% | 9.721% | 14.58% | 23.8465 | 26.3567
Ta, ppm | 1.6552 | 0.24899 | 1.1572 | 2.1531 | 0.90819 | 2.4021 | 15.04% | 30.09% | 45.13% | 1.5724 | 1.7379
Tb, ppm | 0.3716 | 0.02361 | 0.32438 | 0.41882 | 0.30077 | 0.44243 | 6.354% | 12.71% | 19.06% | 0.35302 | 0.39018
Th, ppm | 9.9911 | 0.89358 | 8.204 | 11.7783 | 7.3104 | 12.6719 | 8.944% | 17.89% | 26.83% | 9.4916 | 10.4907
Ti, wt.% | 1.0765 | 0.039826 | 0.99683 | 1.1561 | 0.95701 | 1.196 | 3.7% | 7.399% | 11.1% | 1.0227 | 1.1303
Tl, ppm | 0.2017 | 0.01654 | 0.16863 | 0.23478 | 0.15209 | 0.25132 | 8.2% | 16.4% | 24.6% | 0.19162 | 0.21179
Tm, ppm | 0.2024 | 0.029777 | 0.14285 | 0.26195 | 0.11307 | 0.29173 | 14.71% | 29.42% | 44.14% | 0.19228 | 0.21252
U, ppm | 2.0888 | 0.15049 | 1.7878 | 2.3897 | 1.6373 | 2.5402 | 7.205% | 14.41% | 21.61% | 1.9843 | 2.1932
V, ppm | 252.5181 | 8.3831 | 235.7519 | 269.2842 | 227.3688 | 277.6673 | 3.32% | 6.64% | 9.959% | 239.8922 | 265.144
W, ppm | 1.2688 | 0.1421 | 0.98461 | 1.553 | 0.84251 | 1.6951 | 11.2% | 22.4% | 33.6% | 1.2054 | 1.3322
Y, ppm | 10.8951 | 1.2381 | 8.4188 | 13.3714 | 7.1807 | 14.6095 | 11.36% | 22.73% | 34.09% | 10.3503 | 11.4398
Yb, ppm | 1.2476 | 0.13838 | 0.97083 | 1.5244 | 0.83245 | 1.6628 | 11.09% | 22.18% | 33.28% | 1.1852 | 1.31
Zn, ppm | 35.2584 | 2.2603 | 30.7377 | 39.7791 | 28.4774 | 42.0394 | 6.411% | 12.82% | 19.23% | 33.4955 | 37.0213
Zr, ppm | 171.942 | 11.2396 | 149.4628 | 194.4211 | 138.2232 | 205.6607 | 6.537% | 13.07% | 19.61% | 163.3449 | 180.5391
Aqua Regia Digestion (sample weights 0.15-50g)
Al, wt.% | 4.8054 | 0.6234 | 3.5586 | 6.0522 | 2.9352 | 6.6756 | 12.97% | 25.95% | 38.92% | 4.5651 | 5.0456
Ba, ppm | 157.9904 | 6.538 | 144.9145 | 171.0663 | 138.3765 | 177.6043 | 4.138% | 8.276% | 12.41% | 150.0909 | 165.8899
Be, ppm | 0.98214 | 0.090358 | 0.80143 | 1.1629 | 0.71107 | 1.2532 | 9.2% | 18.4% | 27.6% | 0.93303 | 1.0312
Bi, ppm | 0.16801 | 0.013975 | 0.14006 | 0.19596 | 0.12608 | 0.20993 | 8.318% | 16.64% | 24.95% | 0.15961 | 0.17641
Ca, wt.% | 0.075053 | 0.0064004 | 0.062253 | 0.087854 | 0.055852 | 0.094255 | 8.528% | 17.06% | 25.58% | 0.071301 | 0.078806
Ce, ppm | 22.3091 | 2.2539 | 17.8014 | 26.8169 | 15.5475 | 29.0707 | 10.1% | 20.21% | 30.31% | 21.1937 | 23.4246
Co, ppm | 39.2404 | 3.4399 | 32.3607 | 46.1202 | 28.9208 | 49.5601 | 8.766% | 17.53% | 26.3% | 37.2784 | 41.2025
Cr, ppm | 341.1928 | 25.3831 | 290.4265 | 391.9591 | 265.0434 | 417.3422 | 7.44% | 14.88% | 22.32% | 324.1332 | 358.2524
Cs, ppm | 1.876 | 0.2659 | 1.3442 | 2.4078 | 1.0783 | 2.6737 | 14.17% | 28.35% | 42.52% | 1.7822 | 1.9698
Cu, ppm | 335.8292 | 16.2292 | 303.3707 | 368.2876 | 287.1415 | 384.5169 | 4.833% | 9.665% | 14.5% | 319.0377 | 352.6206
Dy, ppm | 1.4915 | 0.11343 | 1.2646 | 1.7183 | 1.1512 | 1.8317 | 7.605% | 15.21% | 22.82% | 1.4169 | 1.566
Er, ppm | 0.78286 | 0.063233 | 0.65639 | 0.90932 | 0.59316 | 0.97256 | 8.077% | 16.15% | 24.23% | 0.74371 | 0.822
Eu, ppm | 0.49195 | 0.060638 | 0.37067 | 0.61322 | 0.31003 | 0.67386 | 12.33% | 24.65% | 36.98% | 0.46735 | 0.51654
Fe, wt.% | 13.6892 | 0.55959 | 12.57 | 14.8084 | 12.0105 | 15.368 | 4.088% | 8.176% | 12.26% | 13.0048 | 14.3737
Ga, ppm | 20.3281 | 1.9143 | 16.4995 | 24.1566 | 14.5852 | 26.0709 | 9.417% | 18.83% | 28.25% | 19.3117 | 21.3445
Gd, ppm | 1.7022 | 0.15918 | 1.3838 | 2.0205 | 1.2246 | 2.1797 | 9.352% | 18.7% | 28.05% | 1.6171 | 1.7873
Ge, ppm | 0.11746 | 0.017182 | 0.083092 | 0.15182 | 0.065911 | 0.169 | 14.63% | 29.26% | 43.88% | 0.11158 | 0.12333
Hf, ppm | 0.92755 | 0.16557 | 0.59642 | 1.2587 | 0.43085 | 1.4243 | 17.85% | 35.7% | 53.55% | 0.88117 | 0.97393
Hg, ppm | 0.030726 | 0.0050053 | 0.020716 | 0.040737 | 0.01571 | 0.045742 | 16.29% | 32.58% | 48.87% | 0.02919 | 0.032263
Ho, ppm | 0.27807 | 0.022204 | 0.23366 | 0.32248 | 0.21146 | 0.34468 | 7.985% | 15.97% | 23.96% | 0.26417 | 0.29197
In, ppm | 0.087199 | 0.0064147 | 0.07437 | 0.10003 | 0.067955 | 0.10644 | 7.356% | 14.71% | 22.07% | 0.082839 | 0.091559
K, wt.% | 0.082108 | 0.014102 | 0.053904 | 0.11031 | 0.039802 | 0.12441 | 17.18% | 34.35% | 51.53% | 0.078002 | 0.086213
La, ppm | 10.7474 | 0.74942 | 9.2486 | 12.2463 | 8.4992 | 12.9957 | 6.973% | 13.95% | 20.92% | 10.2101 | 11.2848
Lu, ppm | 0.097035 | 0.0068265 | 0.083382 | 0.11069 | 0.076556 | 0.11751 | 7.035% | 14.07% | 21.11% | 0.092183 | 0.10189
Mg, wt.% | 0.15205 | 0.020926 | 0.1102 | 0.1939 | 0.089272 | 0.21483 | 13.76% | 27.52% | 41.29% | 0.14445 | 0.15965
Mn, wt.% | 0.014642 | 0.0016359 | 0.01137 | 0.017914 | 0.0097344 | 0.01955 | 11.17% | 22.35% | 33.52% | 0.01391 | 0.015374
Mo, ppm | 1.1928 | 0.22622 | 0.74032 | 1.6452 | 0.5141 | 1.8714 | 18.97% | 37.93% | 56.9% | 1.1331 | 1.2524
Na, wt.% | 0.032282 | 0.0037512 | 0.02478 | 0.039784 | 0.021029 | 0.043536 | 11.62% | 23.24% | 34.86% | 0.030668 | 0.033896
Nd, ppm | 10.0694 | 1.0122 | 8.045 | 12.0939 | 7.0327 | 13.1062 | 10.05% | 20.11% | 30.16% | 9.566 | 10.5729
Ni, ppm | 191.8644 | 13.6796 | 164.5052 | 219.2235 | 150.8257 | 232.9031 | 7.13% | 14.26% | 21.39% | 182.2712 | 201.4576
P, wt.% | 0.022145 | 0.0014259 | 0.019294 | 0.024997 | 0.017868 | 0.026423 | 6.439% | 12.88% | 19.32% | 0.021038 | 0.023253
Pb, ppm | 12.3904 | 0.59958 | 11.1913 | 13.5896 | 10.5917 | 14.1892 | 4.839% | 9.678% | 14.52% | 11.7709 | 13.01
Pd, ppb | 39.9213 | 2.2634 | 35.3944 | 44.4481 | 33.131 | 46.7115 | 5.67% | 11.34% | 17.01% | 37.9252 | 41.9173
Pr, ppm | 2.6333 | 0.3665 | 1.9003 | 3.3663 | 1.5338 | 3.7328 | 13.92% | 27.84% | 41.75% | 2.5016 | 2.7649
Pt, ppb | 36.653 | 2.012 | 32.6289 | 40.677 | 30.6169 | 42.6891 | 5.489% | 10.98% | 16.47% | 34.8203 | 38.4856
Rb, ppm | 14.3727 | 1.0771 | 12.2184 | 16.527 | 11.1413 | 17.6041 | 7.494% | 14.99% | 22.48% | 13.6541 | 15.0914
S, wt.% | 0.026561 | 0.0040637 | 0.018434 | 0.034689 | 0.01437 | 0.038752 | 15.3% | 30.6% | 45.9% | 0.025233 | 0.027889
Sc, ppm | 31.4172 | 2.2581 | 26.9009 | 35.9335 | 24.6428 | 38.1916 | 7.188% | 14.38% | 21.56% | 29.8463 | 32.988
Sm, ppm | 1.9094 | 0.24512 | 1.4192 | 2.3996 | 1.1741 | 2.6448 | 12.84% | 25.67% | 38.51% | 1.8139 | 2.0049
Sn, ppm | 1.9655 | 0.20629 | 1.5529 | 2.378 | 1.3466 | 2.5843 | 10.5% | 20.99% | 31.49% | 1.8672 | 2.0637
Sr, ppm | 13.1565 | 1.1115 | 10.9335 | 15.3795 | 9.822 | 16.491 | 8.448% | 16.9% | 25.34% | 12.4987 | 13.8143
Tb, ppm | 0.24841 | 0.014473 | 0.21946 | 0.27735 | 0.20499 | 0.29183 | 5.826% | 11.65% | 17.48% | 0.23599 | 0.26083
Th, ppm | 7.6691 | 0.6611 | 6.3469 | 8.9913 | 5.6858 | 9.6524 | 8.62% | 17.24% | 25.86% | 7.2857 | 8.0526
Ti, wt.% | 0.097419 | 0.019964 | 0.057491 | 0.13735 | 0.037526 | 0.15731 | 20.49% | 40.99% | 61.48% | 0.092548 | 0.10229
Tl, ppm | 0.11538 | 0.01209 | 0.091196 | 0.13956 | 0.079106 | 0.15165 | 10.48% | 20.96% | 31.44% | 0.10961 | 0.12114
Tm, ppm | 0.10731 | 0.0081412 | 0.091031 | 0.1236 | 0.082889 | 0.13174 | 7.586% | 15.17% | 22.76% | 0.10195 | 0.11268
U, ppm | 1.0877 | 0.090751 | 0.90624 | 1.2692 | 0.81549 | 1.36 | 8.343% | 16.69% | 25.03% | 1.0334 | 1.1421
V, ppm | 217.0678 | 10.5094 | 196.049 | 238.0865 | 185.5396 | 248.5959 | 4.842% | 9.683% | 14.52% | 206.2144 | 227.9211
Y, ppm | 6.7424 | 0.33559 | 6.0712 | 7.4135 | 5.7356 | 7.7491 | 4.977% | 9.955% | 14.93% | 6.4052 | 7.0795
Yb, ppm | 0.68935 | 0.045207 | 0.59894 | 0.77977 | 0.55373 | 0.82497 | 6.558% | 13.12% | 19.67% | 0.65489 | 0.72382
Zn, ppm | 22.2014 | 3.6624 | 14.8765 | 29.5262 | 11.2141 | 33.1887 | 16.5% | 32.99% | 49.49% | 21.0913 | 23.3114
Zr, ppm | 30.0433 | 4.2754 | 21.4925 | 38.5941 | 17.2171 | 42.8695 | 14.23% | 28.46% | 42.69% | 28.5412 | 31.5455
Infrared Combustion
C, wt.% | 0.31845 | 0.020695 | 0.27706 | 0.35984 | 0.25637 | 0.38053 | 6.498% | 13% | 19.5% | 0.30253 | 0.33437

## Fire Assay
| Analytical results for Au in OREAS 45f (Certified Value 19.3 ppb) | [31 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 04 | 05 | 07 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 37 | 38 | 39 | ppb
|  |  | INAA | FA*AAS | FA*MS | FA*OES | FA*MS | FA*AAS | FA*AAS | FA*MS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*OES | 1
|  |  | 0.085g | 50g | 50g | 30g | 30g | 50g | 40g | 25g | 30g | 25g | 40g | 30g | 40g | 30g | 30g | 25g | 50g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 25g | 30g | 1
| 1 | 1 | 34.2828 | 17 | < 50 | 25 | 17 | 22 | 21 | 18 | 15 | 22 | 5 | 20 | 20 | 20 | 17 | 20 | 22.8975 | 18.5 | 19.6192 | 20 | 16 | 20 | 17 | 19 | 20 | 18 | 20 | 20 | 19.3333 | 18 | 1
| 1 | 2 | 22.5652 | 18 | < 50 | 28 | 16 | 21 | 20 | 18 | 16 | 21 | 5 | 19 | 19 | 19 | 18 | 19 | 21.5 | 21.7 | 19.8505 | 21 | 19 | 20 | 16 | 18 | 20 | 25 | 10 | 21 | 19.6667 | 17 | 30
| 1 | 3 | 14.8201 | 16 | < 50 | 24 | 16 | 21 | 19 | 19 | 17 | 22 | 5 | 22 | 20 | 21 | 16 | 20 | 23.4 | 18.7 | 19.6844 | 20 | 18 | 21 | 17 | 19 | 20 | 21 | 10 | 21 | 19.3333 | 20 | 16
| 1 | 4 | 16.1511 | 17 | < 50 | 24 | 14 | 20 | 20 | 18 | 16 | 22 | 4 | 21 | 20 | 20 | 18 | 20 | 20.7475 | 18.7 | 19.6192 | 20 | 17 | 20 | 17 | 20 | 20 | 20 | 10 | 21 | 21.3333 | 20 | 19.2844
| 1 | 5 | 21.082 | 18 | < 50 | 30 | 16 | 20 | 20 | 18 | 17 | 22 | 5 | 21 | 19 | 20 | 17 | 19 | 24.64 | 18.4 | 19.5183 | 21 | 17 | 20 | 18 | 20 | 20 | 16 | 10 | 21 | 19.6667 | 19 | 7
| 1 | 6 | 13.7379 | 18 | < 50 | 24 | 16 | 19 | 19 | 18 | 17 | 23 | 6 | 20 | 19 | 19 | 18 | 20 | 22.22 | 18.4 | 19.4167 | 18 | 17 | 20 | 16 | 19 | 20 | 17 | 20 | 21 | 19.6667 | 19
|  | 7 | 00: 27.2709
|  | 8 | 00: 18.446
|  | 9 | 00: 19.8429
|  | 10 | 00: 17.172
|  | 11 | 00: 24.562
|  | 12 | 00: 18.1893
|  | 13 | 00: 16.5432
|  | 14 | 00: 17.8906
|  | 15 | 00: 32.3595
|  | 16 | 00: 15.6715
|  | 17 | 00: 14.9831
|  | 18 | 00: 22.9522
|  | 19 | 00: 19.7368
|  | 20 | 00: 14.3864
| Mean |  | 20.1323 | 17.3333 |  | 25.8333 | 15.8333 | 20.5 | 19.8333 | 18.1667 | 16.3333 | 22 | 5 | 20.5 | 19.5 | 19.8333 | 17.3333 | 19.6667 | 22.5675 | 19.0667 | 19.618 | 20 | 17.3333 | 20.1667 | 16.8333 | 19.1667 | 20 | 19.5 | 13.3333 | 20.8333 | 19.8333 | 18.8333
| Median |  | 18.3176 | 17.5 |  | 24.5 | 16 | 20.5 | 20 | 18 | 16.5 | 22 | 5 | 20.5 | 19.5 | 20 | 17.5 | 20 | 22.5587 | 18.6 | 19.6192 | 20 | 17 | 20 | 17 | 19 | 20 | 19 | 10 | 21 | 19.6667 | 19
| Std Dev. |  | 5.772 | 0.8165 |  | 2.5626 | 0.98319 | 1.0488 | 0.75277 | 0.40825 | 0.8165 | 0.63246 | 0.63246 | 1.0488 | 0.54772 | 0.75277 | 0.8165 | 0.5164 | 1.3907 | 1.2972 | 0.14761 | 1.0954 | 1.0328 | 0.40825 | 0.75277 | 0.75277 | 0 | 3.2711 | 5.164 | 0.40825 | 0.75277 | 1.169
| Rel.Std.Dev. |  | 28.67% | 4.711% |  | 9.92% | 6.21% | 5.116% | 3.795% | 2.247% | 4.999% | 2.875% | 12.65% | 5.116% | 2.809% | 3.795% | 4.711% | 2.626% | 6.162% | 6.803% | 0.7524% | 5.477% | 5.958% | 2.024% | 4.472% | 3.928% | 0% | 16.77% | 38.73% | 1.96% | 3.795% | 6.207%
| PDM3 |  | 4.397% | -10.12% |  | 33.96% | -17.9% | 6.304% | 2.847% | -5.796% | -15.3% | 14.08% | -74.07% | 6.304% | 1.118% | 2.847% | -10.12% | 1.982% | 17.02% | -1.129% | 1.73% | 3.711% | -10.12% | 4.575% | -12.71% | -0.6105% | 3.711% | 1.118% | -30.86% | 8.032% | 2.847% | -2.339%
| Z-Score (Absolute) |  | NA | 1.82 | 4.21 | 4.87 | 2.99 | 0.67 | 0.15 | 1.16 | 2.6 | 1.85 | 11.51 | 0.67 | 0.11 | 0.15 | 1.82 | 0.02 | 2.3 | 0.45 | 0.02 | 0.28 | 1.82 | 0.41 | 2.21 | 0.37 | 0.28 | 0.11 | NA | 0.94 | 0.15 | 0.64
Results from laboratory 35 were removed due to their 10 ppb reading resolution.
| Analytical results for Pd in OREAS 45f (Certified Value 56.6 ppb) | [31 empty] | Certified
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 12 | 13 | 14 | 15 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 38 | 39 | [5 empty] | ppb
|  |  | FA*AAS | FA*MS | FA*OES | FA*MS | FA*AAS | FA*AAS | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | [5 empty] | 1
|  |  | 50g | 50g | 30g | 30g | 50g | 40g | 25g | 30g | 25g | 40g | 30g | 30g | 25g | 50g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | [5 empty] | 1
| 1 | 1 | 53 | 53.8848 | 63 | 61.1 | 47 | 57 | 54 | 57 | 51 | 56 | 62 | 55 | 58 | 47 | 59.6 | 53.854 | 56 | 54 | 60 | 57 | 56 | 57.4 | 63 | 57 | 50 | [5 empty] | 1
| 1 | 2 | 51 | 53.5173 | 62 | 60.7 | 47 | 60 | 55 | 58 | 51 | 55 | 59 | 56 | 61.2 | 49.9 | 60.3 | 50.55 | 51 | 55 | 59 | 58 | 55 | 57.1 | 64 | 57.6667 | 47 | [5 empty] | 17
| 1 | 3 | 52 | 53.2863 | 59 | 60.4 | 49 | 58 | 56 | 58 | 49 | 56 | 65 | 51 | 59.4 | 52.4536 | 63.2 | 50.335 | 61 | 55 | 60 | 58 | 57 | 55.6 | 63 | 56.3333 | 53 | [5 empty] | 16
| 1 | 4 | 49 | 53.604 | 60 | 60.3 | 50 | 58 | 54 | 58 | 53 | 56 | 60 | 56 | 56.8 | 53.4 | 59.1 | 50.522 | 56 | 58 | 59 | 60 | 57 | 57.6 | 64 | 56.6667 | 48 | [5 empty] | 56.562
| 1 | 5 | 53 | 53.8887 | 59 | 61.2 | 47 | 60 | 57 | 59 | 50 | 56 | 64 | 56 | 60.6 | 48.7 | 62 | 50.304 | 60 | 54 | 59 | 59 | 59 | 55.4 | 56 | 56.6667 | 54 | [5 empty] | 8
| 1 | 6 | 50 | 53.908 | 61 | 60.2 | 48 | 58 | 57 | 57 | 52 | 56 | 61 | 56 | 58 | 48 | 61.1 | 50.772 | 62 | 56 | 58 | 58 | 58 | 56.1 | 62 | 56.6667 | 51
| Mean |  | 51.3333 | 53.6815 | 60.6667 | 60.65 | 48 | 58.5 | 55.5 | 57.8333 | 51 | 55.8333 | 61.8333 | 55 | 59 | 49.9089 | 60.8833 | 51.0562 | 57.6667 | 55.3333 | 59.1667 | 58.3333 | 57 | 56.5333 | 62 | 56.8333 | 50.5
| Median |  | 51.5 | 53.7444 | 60.5 | 60.55 | 47.5 | 58 | 55.5 | 58 | 51 | 56 | 61.5 | 56 | 58.7 | 49.3 | 60.7 | 50.536 | 58 | 55 | 59 | 58 | 57 | 56.6 | 63 | 56.6667 | 50.5
| Std Dev. |  | 1.633 | 0.25484 | 1.633 | 0.42308 | 1.2649 | 1.2247 | 1.3784 | 0.75277 | 1.4142 | 0.40825 | 2.3166 | 2 | 1.6971 | 2.5389 | 1.5381 | 1.381 | 4.1312 | 1.5055 | 0.75277 | 1.0328 | 1.4142 | 0.95429 | 3.0332 | 0.45947 | 2.7386
| Rel.Std.Dev. |  | 3.181% | 0.4747% | 2.692% | 0.6976% | 2.635% | 2.094% | 2.484% | 1.302% | 2.773% | 0.7312% | 3.747% | 3.636% | 2.876% | 5.087% | 2.526% | 2.705% | 7.164% | 2.721% | 1.272% | 1.771% | 2.481% | 1.688% | 4.892% | 0.8084% | 5.423%
| PDM3 |  | -9.244% | -5.093% | 7.257% | 7.228% | -15.14% | 3.426% | -1.878% | 2.248% | -9.833% | -1.288% | 9.32% | -2.762% | 4.31% | -11.76% | 7.64% | -9.734% | 1.953% | -2.172% | 4.605% | 3.132% | 0.7744% | -0.05061% | 9.614% | 0.4798% | -10.72%
| Z-Score (Absolute) |  | 1.59 | 0.91 | 1.11 | 1.1 | 2.55 | 0.48 | 0.39 | 0.29 | 1.69 | 0.29 | 1.44 | 0.53 | 0.63 | 2 | 1.17 | 1.67 | 0.24 | 0.43 | 0.67 | 0.43 | 0.05 | 0.09 | 1.49 | 0 | 1.83
| Analytical results for Pt in OREAS 45f (Certified Value 38.1 ppb) | [31 empty] | Certified
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 38 | 39 |  |  |  |  | ppb
|  |  | FA*AAS | FA*MS | FA*OES | FA*MS | FA*AAS | FA*AAS | FA*MS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*OES | FA*OES | FA*OES |  |  |  |  | 1
|  |  | 50g | 50g | 30g | 30g | 50g | 40g | 25g | 30g | 25g | 40g | 40g | 30g | 30g | 25g | 50g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 30g |  |  |  |  | 1
| 1 | 1 | 38 | < 50 | 42 | 42.2 | 35 | 38 | 38 | 38 | 30 |  | 40 | 38 | 38 | 39.5 | 33.8 | 42.3 | 31.035 | 35 | 38 | 39 | 40 | 37 | 38.1 | 43 | 40.3333 | 37 |  |  |  |  | 1
| 1 | 2 | 37 | < 50 | 38 | 41.9 | 32 | 38 | 39 | 38 | 29 |  | 40 | 37 | 34 | 40.2 | 36.3 | 43.9 | 28.105 | 32 | 38 | 38 | 38 | 38 | 37.4 | 46 | 39 | 36 |  |  |  |  | 10
| 1 | 3 | 37 | < 50 | 38 | 41.6 | 33 | 38 | 40 | 37 | 31 |  | 40 | 38 | 32 | 38.8 | 35.4 | 43.7 | 25.307 | 38 | 37 | 40 | 38 | 39 | 36.5 | 36 | 38.6667 | 38 |  |  |  |  | 16
| 1 | 4 | 35 | < 50 | 40 | 40.7 | 33 | 37 | 38 | 39 | 28 |  | 40 | 36 | 36 | 39 | 35.7389 | 40.5 | 29.509 | 35 | 39 | 40 | 41 | 38 | 37.3 | 38 | 40 | 32 |  |  |  |  | 38.079
| 1 | 5 | 38 | < 50 | 37 | 41.4 | 35 | 38 | 38 | 39 | 31 |  | 40 | 44 | 33 | 39.5 | 33.8 | 42.6 | 23.239 | 36 | 36 | 40 | 41 | 38 | 35.9 | 38 | 39.6667 | 42 |  |  |  |  | 9
| 1 | 6 | 36 | < 50 | 40 | 41.9 | 32 | 38 | 39 | 38 | 30 |  | 40 | 39 | 38 | 38.8 | 33.9853 | 41.6 | 27.943 | 39 | 39 | 39 | 37 | 39 | 36.7 | 39 | 39.3333 | 38
| Mean |  | 36.8333 |  | 39.1667 | 41.6167 | 33.3333 | 37.8333 | 38.6667 | 38.1667 | 29.8333 |  | 40 | 38.6667 | 35.1667 | 39.3 | 34.8374 | 42.4333 | 27.523 | 35.8333 | 37.8333 | 39.3333 | 39.1667 | 38.1667 | 36.9833 | 40 | 39.5 | 37.1667
| Median |  | 37 |  | 39 | 41.75 | 33 | 38 | 38.5 | 38 | 30 |  | 40 | 38 | 35 | 39.25 | 34.6926 | 42.45 | 28.024 | 35.5 | 38 | 39.5 | 39 | 38 | 37 | 38.5 | 39.5 | 37.5
| Std Dev. |  | 1.169 |  | 1.8348 | 0.52694 | 1.3663 | 0.40825 | 0.8165 | 0.75277 | 1.169 |  | 0 | 2.8048 | 2.5626 | 0.54406 | 1.1088 | 1.2832 | 2.8304 | 2.4833 | 1.169 | 0.8165 | 1.7224 | 0.75277 | 0.77567 | 3.7417 | 0.62361 | 3.2506
| Rel.Std.Dev. |  | 3.174% |  | 4.685% | 1.266% | 4.099% | 1.079% | 2.112% | 1.972% | 3.919% |  | 0% | 7.254% | 7.287% | 1.384% | 3.183% | 3.024% | 10.28% | 6.93% | 3.09% | 2.076% | 4.398% | 1.972% | 2.097% | 9.354% | 1.579% | 8.746%
| PDM3 |  | -3.271% |  | 2.856% | 9.29% | -12.46% | -0.6451% | 1.543% | 0.2303% | -21.65% |  | 5.045% | 1.543% | -7.648% | 3.207% | -8.513% | 11.44% | -27.72% | -5.897% | -0.6451% | 3.294% | 2.856% | 0.2303% | -2.877% | 5.045% | 3.732% | -2.396%
| Z-Score (Absolute) |  | 0.67 | 6.66 | 0.51 | 1.74 | 2.44 | 0.17 | 0.25 | 0 | 4.21 | NA | 0.93 | 0.25 | 1.52 | 0.57 | 1.68 | 2.16 | 5.38 | 1.18 | 0.17 | 0.59 | 0.51 | 0 | 0.6 | 0.93 | 0.67 | 0.51
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C50: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C68: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fire Assay (NiS)
| Analytical results for Au in OREAS 45f (Indicative Value 17.7 ppb) |  |  | Indicative
Au | Round | Replicate | Lab | 1
| No | No | 40 | ppb
|  |  | NiS*MS | 1
|  |  | 25g | 1
| 1 | 1 | 17.7823 | 1
| 1 | 2 | 17.9594 | 30
| 1 | 3 | 18.2628 | 16
| 1 | 4 | 17.0525 | 17.6667
| 1 | 5 | 17.5704 | 36
| 1 | 6 | 17.3729
| Mean |  | 17.6667
| Median |  | 17.6764
| Std Dev. |  | 0.43061
| Rel.Std.Dev. |  | 2.437%
| PDM3 |  | 1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ir in OREAS 45f (Indicative Value 1.45 ppb) |  |  | Indicative
Ir | Round | Replicate | Lab | 1
| No | No | 40 | ppb
|  |  | NiS*MS | 2
|  |  | 25g | 2
| 1 | 1 | 1.4591 | 1
| 1 | 2 | 1.4387 | 26
| 1 | 3 | 1.4723 | 16
| 1 | 4 | 1.4298 | 1.4462
| 1 | 5 | 1.4546 | 37
| 1 | 6 | 1.4224
| Mean |  | 1.4462
| Median |  | 1.4467
| Std Dev. |  | 0.019062
| Rel.Std.Dev. |  | 1.318%
| PDM3 |  | -2.331e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Pd in OREAS 45f (Indicative Value 56 ppb) |  |  | Indicative
Pd | Round | Replicate | Lab | 1
| No | No | 40 | ppb
|  |  | NiS*MS | 1
|  |  | 25g | 1
| 1 | 1 | 56.4477 | 1
| 1 | 2 | 55.5395 | 17
| 1 | 3 | 57.222 | 16
| 1 | 4 | 55.2731 | 55.9819
| 1 | 5 | 56.5947 | 36
| 1 | 6 | 54.814
| Mean |  | 55.9819
| Median |  | 55.9936
| Std Dev. |  | 0.91569
| Rel.Std.Dev. |  | 1.636%
| PDM3 |  | -5.551e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 45f (Indicative Value 36.6 ppb) |  |  | Indicative
Pt | Round | Replicate | Lab | 1
| No | No | 40 | ppb
|  |  | NiS*MS | 1
|  |  | 25g | 1
| 1 | 1 | 36.6967 | 1
| 1 | 2 | 36.4626 | 10
| 1 | 3 | 37.3088 | 16
| 1 | 4 | 35.934 | 36.588
| 1 | 5 | 36.9705 | 37
| 1 | 6 | 36.1556
| Mean |  | 36.588
| Median |  | 36.5797
| Std Dev. |  | 0.51149
| Rel.Std.Dev. |  | 1.398%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rh in OREAS 45f (Indicative Value 3.32 ppb) |  |  | Indicative
Rh | Round | Replicate | Lab | 1
| No | No | 40 | ppb
|  |  | NiS*MS | 2
|  |  | 25g | 2
| 1 | 1 | 3.3314 | 1
| 1 | 2 | 3.3084 | 12
| 1 | 3 | 3.3744 | 16
| 1 | 4 | 3.3119 | 3.3187
| 1 | 5 | 3.3239 | 36
| 1 | 6 | 3.2625
| Mean |  | 3.3187
| Median |  | 3.3179
| Std Dev. |  | 0.036358
| Rel.Std.Dev. |  | 1.096%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 45f (Indicative Value 3.31 ppb) |  |  | Indicative
Ru | Round | Replicate | Lab | 1
| No | No | 40 | ppb
|  |  | NiS*MS | 2
|  |  | 25g | 2
| 1 | 1 | 3.4109 | 1
| 1 | 2 | 3.1931 | 13
| 1 | 3 | 3.3812 | 16
| 1 | 4 | 3.264 | 3.3065
| 1 | 5 | 3.3113 | 37
| 1 | 6 | 3.2785
| Mean |  | 3.3065
| Median |  | 3.2949
| Std Dev. |  | 0.07993
| Rel.Std.Dev. |  | 2.417%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 45f (Certified Value 18 ppb) | [21 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 27 | 28 | 29 | 31 | 33 | 34 | 35 | 37 | ppb
|  |  | INAA | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | 1
|  |  | 0.085g | 25g | 30g | 40g | 25g | 25g | 50g | 40g | 15g | 15g | 25g | 25g | 50g | 25g | 25g | 25g | 25g | 50g | 25g | 50g | 1
| 1 | 1 | 34.2828 | 21 | < 30 | 18 | 17 | 17.1 | 4 | 20 | 17.9 | 15 | 17 | 18.7528 | 20 | 18.7 | 17.5 | 16.5 | 17 | 16 | 17.8 | 21 | 1
| 1 | 2 | 22.5652 | 19 | < 30 | 18 | 19 | 17.9 | 4 | 18 | 20.9 | 15 | 17 | 18.5908 | 23 | 19 | 17.5 | 16.1 | 19 | 17 | 17.5 | 20 | 30
| 1 | 3 | 14.8201 | 21 | < 30 | 16 | 19 | 17.5 | 5 | 20 | 17.5 | 15 | 20 | 19.7676 | 20 | 18.7 | 18.1 | 16.8 | 16 | N.A. | 17.2 | 20 | 16
| 1 | 4 | 16.1511 | 20 | < 30 | 16 | 19 | 17.9 | 4 | 18 | 20 | 14 | 20 | 19.5641 | 24 | 18.5 | 17.2 | 15.8 | 15 | 18 | 17.9 | 21 | 18.0446
| 1 | 5 | 21.082 | 18 | < 30 | 16 | 17 | 16.7 | 4 | 20 | 19.5 | 16 | 17 | 19.3093 | 24 | 18.4 | 17.5 | 18.8 | 18 | 15 | 17.7 | 21 | 11
| 1 | 6 | 13.7379 | 25 | < 30 | 16 | 19 | 18.2 | 3 | 19 | 19.5 | 15 | 18 | 19.337 | 22 | 18.5 | 18.9 | 17.4 | 19 | 19 | 17.9 | 20
|  | 7 | 27.2709
|  | 8 | 18.446
|  | 9 | 19.8429
|  | 10 | 17.172
|  | 11 | 24.562
|  | 12 | 18.1893
|  | 13 | 16.5432
|  | 14 | 17.8906
|  | 15 | 32.3595
|  | 16 | 15.6715
|  | 17 | 14.9831
|  | 18 | 22.9522
|  | 19 | 19.7368
|  | 20 | 14.3864
| Mean |  | 20.1323 | 20.6667 |  | 16.6667 | 18.3333 | 17.55 | 4 | 19.1667 | 19.2167 | 15 | 18.1667 | 19.2203 | 22.1667 | 18.6333 | 17.7833 | 16.9 | 17.3333 | 17 | 17.6667 | 20.5
| Median |  | 18.3176 | 20.5 |  | 16 | 19 | 17.7 | 4 | 19.5 | 19.5 | 15 | 17.5 | 19.3232 | 22.5 | 18.6 | 17.5 | 16.65 | 17.5 | 17 | 17.75 | 20.5
| Std Dev. |  | 5.772 | 2.4221 |  | 1.0328 | 1.0328 | 0.5648 | 0.63246 | 0.98319 | 1.2875 | 0.63246 | 1.472 | 0.45913 | 1.8348 | 0.21602 | 0.62102 | 1.0844 | 1.633 | 1.5811 | 0.27325 | 0.54772
| Rel.Std.Dev. |  | 28.67% | 11.72% |  | 6.197% | 5.633% | 3.218% | 15.81% | 5.13% | 6.7% | 4.216% | 8.103% | 2.389% | 8.278% | 1.159% | 3.492% | 6.417% | 9.421% | 9.301% | 1.547% | 2.672%
| PDM3 |  | 11.57% | 14.53% |  | -7.636% | 1.6% | -2.741% | -77.83% | 6.218% | 6.495% | -16.87% | 0.6765% | 6.515% | 22.84% | 3.263% | -1.448% | -6.343% | -3.942% | -5.789% | -2.094% | 13.61%
| Z-Score (Absolute) |  | NA | 1.74 | 1.68 | 0.67 | 0.33 | 0.14 | 8.32 | 0.84 | 0.87 | 1.68 | 0.23 | 0.87 | 2.65 | 0.51 | 0 | 0.53 | 0.27 | 0.47 | 0.07 | 1.64
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 45f (Certified Value 20.46 wt.%) | [21 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 20.812 | 20.6 | 20.47 | 20.7487 | 20.5 | 20.43 | 20.38 | 20.8 | 20.3 | 20.5 | 19.7483 | 20.3 | 20.53 | 20.3 | 20.32 | 20.5 | 20.3 | 20.42 | 20.3924 | 20.087 | 1
| 1 | 2 | 20.851 | 20.8 | 20.48 | 20.5778 | 20.58 | 20.46 | 20.25 | 20.8 | 20.4 | 20.59 | 19.9562 | 20.4 | 20.5 | 20.38 | 20.59 | 20.5 | 20.3 | 20.37 | 20.3619 | 20.1058 | #N/A
| 1 | 3 | 20.801 | 20.7 | 20.43 | 20.4896 | 20.61 | 20.42 | 20.33 | 20.8 | 20.4 | 20.55 | 19.9384 | 20.3 | 20.52 | 20.4 | 20.35 | 20.4 | 20.2 | 20.44 | 20.5375 | 20.2075 | 16
| 1 | 4 | 20.985 | 20.7 | 20.47 | 20.6413 | 20.52 | 20.37 | 20.22 | 20.8 | 20.4 | 20.51 | 19.9739 | 20.3 | 20.59 | 20.42 | 20.42 | 20.4 | 20.2 | 20.35 | 19.9906 | 20.2178 | 20.4579
| 1 | 5 | 20.815 | 20.7 | 20.46 | 20.5867 | 20.54 | 20.45 | 20.34 | 20.6 | 20.4 | 20.51 | 19.9357 | 20.2 | 20.36 | 20.36 | 20.48 | 20.4 | 20.4 | 20.43 | 20.2722 | 20.2696 | 13
| 1 | 6 | 20.782 | 20.6 | 20.52 | 20.6159 | 20.54 | 20.53 | 20.22 | 20.8 | 20.3 | 20.5 | 19.9566 | 20.3 | 20.24 | 20.44 | 20.45 | 20.4 | 20.4 | 20.48 | 20.1076 | 20.0278
| Mean |  | 20.841 | 20.6833 | 20.4717 | 20.61 | 20.5483 | 20.4433 | 20.29 | 20.7667 | 20.3667 | 20.5267 | 19.9182 | 20.3 | 20.4567 | 20.3833 | 20.435 | 20.4333 | 20.3 | 20.415 | 20.277 | 20.1526
| Median |  | 20.8135 | 20.7 | 20.47 | 20.6013 | 20.54 | 20.44 | 20.29 | 20.8 | 20.4 | 20.51 | 19.9473 | 20.3 | 20.51 | 20.39 | 20.435 | 20.4 | 20.3 | 20.425 | 20.317 | 20.1567
| Std Dev. |  | 0.074073 | 0.075277 | 0.029269 | 0.085259 | 0.040208 | 0.052789 | 0.068702 | 0.08165 | 0.05164 | 0.036148 | 0.084373 | 0.063246 | 0.13064 | 0.049666 | 0.096902 | 0.05164 | 0.089443 | 0.047645 | 0.19956 | 0.092744
| Rel.Std.Dev. |  | 0.3554% | 0.364% | 0.143% | 0.4137% | 0.1957% | 0.2582% | 0.3386% | 0.3932% | 0.2536% | 0.1761% | 0.4236% | 0.3116% | 0.6386% | 0.2437% | 0.4742% | 0.2527% | 0.4406% | 0.2334% | 0.9842% | 0.4602%
| PDM3 |  | 1.872% | 1.102% | 0.06716% | 0.7434% | 0.4419% | -0.07134% | -0.8208% | 1.509% | -0.4461% | 0.336% | -2.638% | -0.772% | -0.006163% | -0.3646% | -0.1121% | -0.1202% | -0.772% | -0.2098% | -0.8843% | -1.492%
| Z-Score (Absolute) |  | 2.21 | 1.35 | 0.2 | 0.95 | 0.62 | 0.05 | 0.78 | 1.81 | 0.37 | 0.5 | 2.8 | 0.73 | 0.12 | 0.28 | 0 | 0 | 0.73 | 0.1 | 0.85 | 1.53
| Analytical results for As in OREAS 45f (Indicative Value 9.4 ppm) | [21 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 12 | 19 | 33 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [16 empty] | 2
2
| 1 | 1 | 8 | 10 | 10 | < 70 | [16 empty] | 1
| 1 | 2 | 8 | 10 | 10 | < 70 | [16 empty] | 1
| 1 | 3 | < 8 | 10 | 10 | < 70 | [16 empty] | 16
| 1 | 4 | 8 | 10 | 10 | < 70 | [16 empty] | 9.4017
| 1 | 5 | 15 | 10 | 10 | < 70 | [16 empty] | 7
| 1 | 6 | 8 | 10 | 10 | < 70
| Mean |  | 9.4 | 10 | 10
| Median |  | 8 | 10 | 10
| Std Dev. |  | 3.1305 | 0 | 0
| Rel.Std.Dev. |  | 33.3% | 0% | 0%
| PDM3 |  | -0.01846% | 6.363% | 6.363%
| Z-Score (Absolute) |  | 1.35 | 0 | 0 | 16.98
| Analytical results for BaO in OREAS 45f (Certified Value 266 ppm) | [21 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 10 | 12 | 14 | 17 | 20 | 23 | 24 | 28 | 29 | 31 | 33 | 35 | 39 | 40 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 300 | 173.1423 | 279 | 200 | < 100 | 110 | 300 | 315.4217 | 200 | 100 | 200 | 300 | 200 | 508.8668 | 259.7838 | [5 empty] | 1
| 1 | 2 | 300 | 162.3098 | 290 | 200 | < 100 | 110 | 300 | 320.8694 | 300 | 100 | 200 | 300 | 200 | 519.8396 | 267.8156 | [5 empty] | #N/A
| 1 | 3 | 270 | 214.5031 | 279 | 200 | < 100 | 220 | 300 | 309.6306 | 300 | 100 | 200 | 300 | 200 | 427.9994 | 258.9604 | [5 empty] | 16
| 1 | 4 | 320 | 199.3733 | 279 | 200 | < 100 | 110 | 300 | 321.6062 | 300 | 100 | 200 | 300 | 200 | 427.5122 | 249.0621 | [5 empty] | 265.7438
| 1 | 5 | 280 | 200.9848 | 246 | 200 | < 100 | 220 | 300 | 313.1399 | 200 | 100 | 200 | 300 | 200 | 427.7778 | 251.7578 | [5 empty] | 14
| 1 | 6 | 390 | 206.5354 | 257 | 200 | < 100 | 220 | 200 | 315.2056 | 200 | 100 | 300 | 300 | 200 | 404.8643 | 238.1218
| Mean |  | 310 | 192.8081 | 271.6667 | 200 |  | 165 | 283.3333 | 315.9789 | 250 | 100 | 216.6667 | 300 | 200 | 452.81 | 254.2502
| Median |  | 300 | 200.1791 | 279 | 200 |  | 165 | 300 | 315.3136 | 250 | 100 | 200 | 300 | 200 | 427.8886 | 255.3591
| Std Dev. |  | 42.8952 | 20.4253 | 16.561 | 0 |  | 60.2495 | 40.8248 | 4.5794 | 54.7723 | 0 | 40.8248 | 0 | 0 | 48.6133 | 10.2964
| Rel.Std.Dev. |  | 13.84% | 10.59% | 6.096% | 0% |  | 36.51% | 14.41% | 1.449% | 21.91% | 0% | 18.84% | 0% | 0% | 10.74% | 4.05%
| PDM3 |  | 16.65% | -27.45% | 2.229% | -24.74% |  | -37.91% | 6.619% | 18.9% | -5.924% | -62.37% | -18.47% | 12.89% | -24.74% | 70.39% | -4.325%
| Z-Score (Absolute) |  | 0.42 | 2.14 | 0.42 | NA | NA | NA | NA | 0.55 | NA | NA | NA | NA | NA | 3.54 | 0.8
Results from laboratories 14, 17, 20, 23, 28, 29, 31, 33 & 35 were removed due to their 100 ppm reading resolution.
| Analytical results for CaO in OREAS 45f (Certified Value 0.139 wt.%) | [21 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.145 | 0.13 | 0.14 | 0.14113 | 0.139 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.16421 | 0.14 | 0.13 | 0.13 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14079 | 0.14239 | 1
| 1 | 2 | 0.145 | 0.13 | 0.14 | 0.13719 | 0.138 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.1657 | 0.14 | 0.13 | 0.13 | 0.14 | 0.14 | 0.16 | 0.14 | 0.14167 | 0.14218 | #N/A
| 1 | 3 | 0.15 | 0.13 | 0.14 | 0.14925 | 0.137 | 0.13 | 0.12 | 0.14 | 0.14 | 0.14 | 0.16293 | 0.14 | 0.13 | 0.13 | 0.14 | 0.14 | 0.15 | 0.14 | 0.146 | 0.15232 | 16
| 1 | 4 | 0.143 | 0.13 | 0.14 | 0.17728 | 0.139 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.16381 | 0.14 | 0.13 | 0.11 | 0.14 | 0.14 | 0.14 | 0.14 | 0.13859 | 0.1462 | 0.1394
| 1 | 5 | 0.148 | 0.13 | 0.14 | 0.146 | 0.138 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.16591 | 0.14 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14282 | 0.14467 | 15
| 1 | 6 | 0.146 | 0.13 | 0.14 | 0.13864 | 0.136 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.16526 | 0.14 | 0.13 | 0.13 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14911 | 0.14512
| Mean |  | 0.14617 | 0.13 | 0.14 | 0.14825 | 0.13783 | 0.13 | 0.12833 | 0.14 | 0.14 | 0.14 | 0.16464 | 0.14 | 0.13 | 0.12667 | 0.14 | 0.14 | 0.14833 | 0.14 | 0.14316 | 0.14548
| Median |  | 0.1455 | 0.13 | 0.14 | 0.14356 | 0.138 | 0.13 | 0.13 | 0.14 | 0.14 | 0.14 | 0.16473 | 0.14 | 0.13 | 0.13 | 0.14 | 0.14 | 0.15 | 0.14 | 0.14224 | 0.1449
| Std Dev. |  | 0.0024833 | 0 | 0 | 0.014928 | 0.001169 | 0 | 0.0040825 | 0 | 0 | 0 | 0.0011768 | 0 | 0 | 0.008165 | 0 | 0 | 0.0075277 | 0 | 0.0038005 | 0.003703
| Rel.Std.Dev. |  | 1.699% | 0% | 0% | 10.07% | 0.8482% | 0% | 3.181% | 0% | 0% | 0% | 0.7148% | 0% | 0% | 6.446% | 0% | 0% | 5.075% | 0% | 2.655% | 2.545%
| PDM3 |  | 4.858% | -6.74% | 0.434% | 6.351% | -1.12% | -6.74% | -7.936% | 0.434% | 0.434% | 0.434% | 18.11% | 0.434% | -6.74% | -9.131% | 0.434% | 0.434% | 6.412% | 0.434% | 2.704% | 4.364%
| Z-Score (Absolute) |  | 0.96 | 1.56 | 0 | 1.29 | 0.34 | 1.56 | 1.82 | 0 | 0 | 0 | 3.84 | 0 | 1.56 | 2.08 | 0 | 0 | 1.3 | 0 | 0.49 | 0.85
| Analytical results for Cl in OREAS 45f (Indicative Value 193 ppm) | [21 empty] | Indicative
Cl | Round | Replicate | Lab | Lab | Lab | Lab | [16 empty] | 1
| No | No | 01 | 12 | 19 | 39 | [16 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [16 empty] | 0
0
| 1 | 1 | 110 | 220 | 240 | 165.7344 | [16 empty] | 1
| 1 | 2 | 160 | 220 | 210 | 196.5856 | [16 empty] | 2
| 1 | 3 | 210 | 210 | 250 | 103.6055 | [16 empty] | 16
| 1 | 4 | 150 | 220 | 230 | 144.8825 | [16 empty] | 193.0409
| 1 | 5 | 160 | 210 | 230 | 165.6829 | [16 empty] | 8
| 1 | 6 | 180 | 220 | 240 | 186.4898
| Mean |  | 161.6667 | 216.6667 | 233.3333 | 160.4968
| Median |  | 160 | 220 | 235 | 165.7087
| Std Dev. |  | 33.116 | 5.164 | 13.6626 | 33.2048
| Rel.Std.Dev. |  | 20.48% | 2.383% | 5.855% | 20.69%
| PDM3 |  | -16.25% | 12.24% | 20.87% | -16.86%
| Z-Score (Absolute) |  | 0.66 | 0.66 | 1.06 | 0.69
| Analytical results for Co in OREAS 45f (Indicative Value 43.7 ppm) | [21 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 10 | 12 | 19 | 33 | 39 | 40 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [13 empty] | 1
1
| 1 | 1 | 63 | 34.2144 | 40 | 40 | < 140 | 103.584 | 48.3881 | [13 empty] | 1
| 1 | 2 | 55 | 32.232 | 30 | 50 | < 140 | 103.4661 | 46.3824 | [13 empty] | 3
| 1 | 3 | 63 | 40.5286 | 40 | 40 | < 140 | 103.6055 | 47.2063 | [13 empty] | 16
| 1 | 4 | 63 | 34.141 | 40 | 30 | < 140 | 103.4875 | 48.3468 | [13 empty] | 43.6747
| 1 | 5 | 63 | 34.8752 | 40 | 50 | < 140 | 103.5518 | 42.5561 | [13 empty] | 9
| 1 | 6 | 55 | 32.746 | 40 | 30 | < 140 | 103.6055 | 43.7335
| Mean |  | 60.3333 | 34.7895 | 38.3333 | 40 |  | 103.5501 | 46.1022
| Median |  | 63 | 34.1777 | 40 | 40 |  | 103.5679 | 46.7944
| Std Dev. |  | 4.1312 | 2.9804 | 4.0825 | 8.9443 |  | 0.06042 | 2.4391
| Rel.Std.Dev. |  | 6.847% | 8.567% | 10.65% | 22.36% |  | 0.05835% | 5.291%
| PDM3 |  | 38.14% | -20.34% | -12.23% | -8.414% |  | 137.1% | 5.558%
| Z-Score (Absolute) |  | 1.79 | 0.86 | 0.49 | 0.32 | NA | 6.29 | 0.32
Results from laboratory 33 were removed due to their 100 ppm reading resolution.
| Analytical results for Cr2O3 in OREAS 45f (Certified Value 707 ppm) | [21 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 35 | 39 | 40 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 800 | 700 | 720 | 841.6 | 702 | 700 | 1000 | 628 | 700 | 700 | < 50 | 700 | 700 | 700 | 700 | 700 | 700 | 639.3233 | 672.3817 |  | 1
| 1 | 2 | 800 | 800 | 710 | 833.6 | 731 | 700 | 900 | 628 | 700 | 700 | < 50 | 800 | 800 | 700 | 700 | 700 | 700 | 659.5332 | 685.7157 |  | #N/A
| 1 | 3 | 680 | 700 | 730 | 753.5 | 687 | 700 | 900 | 702 | 700 | 700 | < 50 | 700 | 700 | 700 | 700 | 800 | 800 | 597.5242 | 684.2668 |  | 16
| 1 | 4 | 700 | 800 | 750 | 783.2 | 702 | 700 | 900 | 658 | 700 | 700 | < 50 | 700 | 700 | 700 | 700 | 700 | 700 | 607.315 | 672.5575 |  | 706.8244
| 1 | 5 | 920 | 700 | 730 | 727.9 | 687 | 700 | 900 | 702 | 700 | 700 | < 50 | 700 | 700 | 700 | 700 | 700 | 700 | 733.4218 | 686.04 |  | 16
| 1 | 6 | 730 | 700 | 740 | 751.3 | 731 | 700 | 1000 | 672 | 700 | 700 | < 50 | 700 | 700 | 700 | 700 | 700 | 800 | 660.4215 | 664.0454
| Mean |  | 771.6667 | 733.3333 | 730 | 781.85 | 706.6667 | 700 | 933.3333 | 665 | 700 | 700 |  | 716.6667 | 716.6667 | 700 | 700 | 716.6667 | 733.3333 | 649.5898 | 677.5012
| Median |  | 765 | 700 | 730 | 768.35 | 702 | 700 | 900 | 665 | 700 | 700 |  | 700 | 700 | 700 | 700 | 700 | 700 | 649.4282 | 678.4122
| Std Dev. |  | 88.1854 | 51.6398 | 14.1421 | 46.6853 | 20.0067 | 0 | 51.6398 | 33.3886 | 0 | 0 |  | 40.8248 | 40.8248 | 0 | 0 | 40.8248 | 51.6398 | 48.6773 | 9.1419
| Rel.Std.Dev. |  | 11.43% | 7.042% | 1.937% | 5.971% | 2.831% | 0% | 5.533% | 5.021% | 0% | 0% |  | 5.696% | 5.696% | 0% | 0% | 5.696% | 7.042% | 7.494% | 1.349%
| PDM3 |  | 9.174% | 3.75% | 3.279% | 10.61% | -0.02232% | -0.9655% | 32.05% | -5.917% | -0.9655% | -0.9655% |  | 1.392% | 1.392% | -0.9655% | -0.9655% | 1.392% | 3.75% | -8.097% | -4.149%
| Z-Score (Absolute) |  | 1.86 | 0.77 | 0.67 | 2.16 | 0 | 0.18 | 6.48 | 1.18 | 0.18 | 0.18 | 19.46 | 0.29 | 0.29 | 0.18 | 0.18 | 0.29 | 0.77 | 1.62 | 0.83
| Analytical results for Cu in OREAS 45f (Certified Value 364 ppm) | [21 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 10 | 12 | 19 | 33 | 39 | 40 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 367 | 378.6547 | 350 | 340 | 400 | 364.6157 | 376.1202 | [13 empty] | 1
| 1 | 2 | 351 | 385.0455 | 350 | 350 | 320 | 364.2007 | 372.9557 | [13 empty] | #N/A
| 1 | 3 | 351 | 379.7731 | 350 | 350 | 400 | 389.5566 | 378.1541 | [13 empty] | 16
| 1 | 4 | 351 | 374.1013 | 360 | 350 | 320 | 364.2761 | 357.3996 | [13 empty] | 364.1755
| 1 | 5 | 344 | 391.1168 | 350 | 350 | 320 | 381.0707 | 374.7235 | [13 empty] | 17
| 1 | 6 | 359 | 376.8174 | 350 | 340 | 320 | 389.5566 | 378.5418
| Mean |  | 353.8333 | 380.9181 | 351.6667 | 346.6667 | 346.6667 | 375.5461 | 372.9825
| Median |  | 351 | 379.2139 | 350 | 350 | 320 | 372.8432 | 375.4219
| Std Dev. |  | 8.0104 | 6.1769 | 4.0825 | 5.164 | 41.3118 | 12.6358 | 7.9165
| Rel.Std.Dev. |  | 2.264% | 1.622% | 1.161% | 1.49% | 11.92% | 3.365% | 2.122%
| PDM3 |  | -2.84% | 4.597% | -3.435% | -4.808% | -4.808% | 3.122% | 2.418%
| Z-Score (Absolute) |  | 0.53 | 0.98 | 0.67 | 0.95 | NA | 0.68 | 0.53
Results from laboratory 33 were removed due to their 100 ppm reading resolution.
| Analytical results for F in OREAS 45f (Indicative Value 1478 ppm) | [21 empty] | Indicative
F | Round | Replicate | Lab | [19 empty] | 1
| No | No | 39 | [19 empty] | ppm
|  |  | BF*XRF | [19 empty] | 0
0
| 1 | 1 | 2019.8881 | [19 empty] | 1
| 1 | 2 | 1034.6611 | [19 empty] | 1
| 1 | 3 | 1844.1774 | [19 empty] | 16
| 1 | 4 | < 1000 | [19 empty] | 1478.0416
| 1 | 5 | 2433.468 | [19 empty] | 7
| 1 | 6 | 1036.0547
| Mean |  | 1673.6499
| Median |  | 1844.1774
| Std Dev. |  | 620.7089
| Rel.Std.Dev. |  | 37.09%
| PDM3 |  | 13.23%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 45f (Certified Value 22.31 wt.%) | [21 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 22.518 | 22.6 | 22.07 | 22.5734 | 22.275 | 22.49 | 22.85 | 21.917 | 22 | 22.19 | 21.7741 | 21.9 | 22.27 | 22.24 | 22.47 | 22.4 | 22.3 | 21.96 | 22.5238 | 22.0417 | 1
| 1 | 2 | 22.558 | 22.5 | 22.14 | 22.542 | 22.117 | 22.45 | 23.05 | 21.96 | 22 | 22.22 | 21.9985 | 22 | 22.31 | 22.32 | 22.78 | 22.3 | 22.5 | 22.13 | 22.7423 | 21.8972 | #N/A
| 1 | 3 | 22.487 | 22.5 | 22.15 | 22.608 | 22.06 | 22.55 | 23.11 | 22.16 | 22.2 | 22.28 | 21.9805 | 22 | 22.56 | 22.31 | 22.59 | 22.4 | 22.3 | 22 | 22.9583 | 22.0475 | 16
| 1 | 4 | 22.551 | 22.5 | 22.16 | 22.5027 | 22.16 | 22.48 | 23.04 | 22.06 | 22.1 | 22.12 | 21.8723 | 22.1 | 22.58 | 22.21 | 22.75 | 22.4 | 22.3 | 22.11 | 22.5514 | 21.7171 | 22.3127
| 1 | 5 | 22.572 | 22.4 | 22.21 | 22.4605 | 22.132 | 22.41 | 22.8 | 22.146 | 22.1 | 22.21 | 21.9968 | 22 | 22.21 | 22.31 | 22.7 | 22.4 | 22.1 | 22 | 22.7538 | 22.0654 | 18
| 1 | 6 | 22.613 | 22.4 | 22.17 | 22.553 | 22.117 | 22.46 | 23.07 | 22.146 | 22.2 | 22.34 | 21.9794 | 22 | 22.14 | 22.27 | 22.79 | 22.4 | 22.1 | 22.07 | 22.6051 | 21.724
| Mean |  | 22.5498 | 22.4833 | 22.15 | 22.5399 | 22.1435 | 22.4733 | 22.9867 | 22.0648 | 22.1 | 22.2267 | 21.9336 | 22 | 22.345 | 22.2767 | 22.68 | 22.3833 | 22.2667 | 22.045 | 22.6891 | 21.9155
| Median |  | 22.5545 | 22.5 | 22.155 | 22.5475 | 22.1245 | 22.47 | 23.045 | 22.103 | 22.1 | 22.215 | 21.98 | 22 | 22.29 | 22.29 | 22.725 | 22.4 | 22.3 | 22.035 | 22.6737 | 21.9695
| Std Dev. |  | 0.043614 | 0.075277 | 0.046043 | 0.052181 | 0.072213 | 0.046762 | 0.12848 | 0.10498 | 0.089443 | 0.075807 | 0.091336 | 0.063246 | 0.1836 | 0.044572 | 0.12617 | 0.040825 | 0.15055 | 0.068337 | 0.163 | 0.1626
| Rel.Std.Dev. |  | 0.1934% | 0.3348% | 0.2079% | 0.2315% | 0.3261% | 0.2081% | 0.5589% | 0.4758% | 0.4047% | 0.3411% | 0.4164% | 0.2875% | 0.8217% | 0.2001% | 0.5563% | 0.1824% | 0.6761% | 0.31% | 0.7184% | 0.7419%
| PDM3 |  | 1.063% | 0.7649% | -0.729% | 1.019% | -0.7582% | 0.7201% | 3.021% | -1.111% | -0.9531% | -0.3854% | -1.699% | -1.401% | 0.1449% | -0.1614% | 1.646% | 0.3167% | -0.2062% | -1.2% | 1.687% | -1.78%
| Z-Score (Absolute) |  | 0.9 | 0.68 | 0.39 | 0.86 | 0.41 | 0.65 | 2.31 | 0.67 | 0.55 | 0.15 | 1.09 | 0.88 | 0.24 | 0.02 | 1.32 | 0.36 | 0.02 | 0.73 | 1.35 | 1.15
| Analytical results for K2O in OREAS 45f (Certified Value 0.284 wt.%) | [21 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.29 | 0.29 | 0.28 | 0.28084 | 0.28 | 0.28 | 0.27 | 0.29 | 0.28 | 0.28 | 0.28004 | 0.29 | 0.28 | 0.28 | 0.28 | 0.29 | 0.3 | 0.29 | 0.30647 | 0.28855 | 1
| 1 | 2 | 0.293 | 0.29 | 0.28 | 0.28175 | 0.279 | 0.28 | 0.27 | 0.29 | 0.28 | 0.28 | 0.28804 | 0.28 | 0.28 | 0.28 | 0.28 | 0.29 | 0.3 | 0.29 | 0.29682 | 0.28669 | #N/A
| 1 | 3 | 0.294 | 0.28 | 0.28 | 0.28298 | 0.283 | 0.28 | 0.28 | 0.29 | 0.28 | 0.29 | 0.27993 | 0.28 | 0.29 | 0.29 | 0.28 | 0.29 | 0.28 | 0.29 | 0.30032 | 0.28773 | 16
| 1 | 4 | 0.294 | 0.28 | 0.28 | 0.2892 | 0.282 | 0.28 | 0.27 | 0.29 | 0.28 | 0.28 | 0.28041 | 0.28 | 0.29 | 0.28 | 0.28 | 0.29 | 0.29 | 0.29 | 0.29067 | 0.28862 | 0.28391
| 1 | 5 | 0.291 | 0.29 | 0.28 | 0.28872 | 0.281 | 0.28 | 0.27 | 0.29 | 0.28 | 0.28 | 0.28858 | 0.28 | 0.28 | 0.28 | 0.28 | 0.28 | 0.28 | 0.29 | 0.3012 | 0.28691 | 19
| 1 | 6 | 0.291 | 0.29 | 0.28 | 0.283 | 0.282 | 0.28 | 0.27 | 0.29 | 0.28 | 0.28 | 0.28581 | 0.28 | 0.28 | 0.28 | 0.28 | 0.29 | 0.29 | 0.29 | 0.29825 | 0.28781
| Mean |  | 0.29217 | 0.28667 | 0.28 | 0.28441 | 0.28117 | 0.28 | 0.27167 | 0.29 | 0.28 | 0.28167 | 0.2838 | 0.28167 | 0.28333 | 0.28167 | 0.28 | 0.28833 | 0.29 | 0.29 | 0.29896 | 0.28772
| Median |  | 0.292 | 0.29 | 0.28 | 0.28299 | 0.2815 | 0.28 | 0.27 | 0.29 | 0.28 | 0.28 | 0.28311 | 0.28 | 0.28 | 0.28 | 0.28 | 0.29 | 0.29 | 0.29 | 0.29929 | 0.28777
| Std Dev. |  | 0.0017224 | 0.005164 | 0 | 0.0036159 | 0.001472 | 0 | 0.0040825 | 0 | 0 | 0.0040825 | 0.0041352 | 0.0040825 | 0.005164 | 0.0040825 | 0 | 0.0040825 | 0.0089443 | 0 | 0.005237 | 0.00080339
| Rel.Std.Dev. |  | 0.5895% | 1.801% | 0% | 1.271% | 0.5235% | 0% | 1.503% | 0% | 0% | 1.449% | 1.457% | 1.449% | 1.823% | 1.449% | 0% | 1.416% | 3.084% | 0% | 1.752% | 0.2792%
| PDM3 |  | 2.909% | 0.9713% | -1.377% | 0.1782% | -0.9659% | -1.377% | -4.312% | 2.145% | -1.377% | -0.7898% | -0.0379% | -0.7898% | -0.2027% | -0.7898% | -1.377% | 1.558% | 2.145% | 2.145% | 5.3% | 1.342%
| Z-Score (Absolute) |  | 1.63 | 0.59 | 0.67 | 0.16 | 0.45 | 0.67 | 2.25 | 1.22 | 0.67 | 0.36 | 0.04 | 0.36 | 0.04 | 0.36 | 0.67 | 0.9 | 1.22 | 1.22 | 2.91 | 0.79
| Analytical results for MgO in OREAS 45f (Certified Value 0.423 wt.%) | [21 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.402 | 0.42 | 0.43 | 0.35088 | 0.407 | 0.43 | 0.48 | 0.42 | 0.42 | 0.43 | 0.41545 | 0.4 | 0.45 | 0.41 | 0.42 | 0.42 | 0.47 | 0.43 | 0.41532 | 0.42159 | 1
| 1 | 2 | 0.419 | 0.43 | 0.44 | 0.34816 | 0.405 | 0.44 | 0.53 | 0.43 | 0.42 | 0.43 | 0.4201 | 0.41 | 0.45 | 0.42 | 0.43 | 0.42 | 0.47 | 0.44 | 0.42411 | 0.42167 | #N/A
| 1 | 3 | 0.424 | 0.43 | 0.42 | 0.34064 | 0.418 | 0.44 | 0.52 | 0.42 | 0.43 | 0.43 | 0.41444 | 0.41 | 0.45 | 0.42 | 0.42 | 0.42 | 0.45 | 0.43 | 0.43809 | 0.43142 | 16
| 1 | 4 | 0.415 | 0.44 | 0.42 | 0.34482 | 0.4 | 0.44 | 0.53 | 0.43 | 0.42 | 0.43 | 0.42422 | 0.4 | 0.46 | 0.42 | 0.42 | 0.41 | 0.46 | 0.43 | 0.43038 | 0.4208 | 0.42251
| 1 | 5 | 0.419 | 0.43 | 0.42 | 0.36362 | 0.4 | 0.44 | 0.51 | 0.42 | 0.43 | 0.43 | 0.42717 | 0.41 | 0.45 | 0.42 | 0.42 | 0.41 | 0.46 | 0.43 | 0.41828 | 0.42421 | 20
| 1 | 6 | 0.416 | 0.44 | 0.42 | 0.35442 | 0.413 | 0.44 | 0.49 | 0.42 | 0.42 | 0.43 | 0.42121 | 0.41 | 0.44 | 0.42 | 0.43 | 0.41 | 0.43 | 0.43 | 0.42675 | 0.42107
| Mean |  | 0.41583 | 0.43167 | 0.425 | 0.35042 | 0.40717 | 0.43833 | 0.51 | 0.42333 | 0.42333 | 0.43 | 0.42043 | 0.40667 | 0.45 | 0.41833 | 0.42333 | 0.415 | 0.45667 | 0.43167 | 0.42549 | 0.42346
| Median |  | 0.4175 | 0.43 | 0.42 | 0.34952 | 0.406 | 0.44 | 0.515 | 0.42 | 0.42 | 0.43 | 0.42065 | 0.41 | 0.45 | 0.42 | 0.42 | 0.415 | 0.46 | 0.43 | 0.42543 | 0.42163
| Std Dev. |  | 0.0074677 | 0.0075277 | 0.0083666 | 0.008032 | 0.0071949 | 0.0040825 | 0.020976 | 0.005164 | 0.005164 | 0 | 0.0049203 | 0.005164 | 0.0063246 | 0.0040825 | 0.005164 | 0.0054772 | 0.015055 | 0.0040825 | 0.0082625 | 0.0040848
| Rel.Std.Dev. |  | 1.796% | 1.744% | 1.969% | 2.292% | 1.767% | 0.9314% | 4.113% | 1.22% | 1.22% | 0% | 1.17% | 1.27% | 1.405% | 0.9759% | 1.22% | 1.32% | 3.297% | 0.9457% | 1.942% | 0.9646%
| PDM3 |  | -1.581% | 2.166% | 0.5884% | -17.06% | -3.632% | 3.744% | 20.71% | 0.1939% | 0.1939% | 1.772% | -0.4931% | -3.751% | 6.505% | -0.9894% | 0.1939% | -1.778% | 8.083% | 2.166% | 0.7043% | 0.2241%
| Z-Score (Absolute) |  | 0.64 | 0.7 | 0.14 | 6.22 | 1.38 | 1.27 | 7.38 | 0.01 | 0.01 | 0.56 | 0.25 | 1.43 | 2.27 | 0.43 | 0.01 | 0.72 | 2.83 | 0.7 | 0.18 | 0.01
| Analytical results for MnO in OREAS 45f (Certified Value 0.031 wt.%) | [21 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.032 | 0.02 | 0.024 | 0.03152 | 0.0323 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033294 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.02889 | 0.033619 | 1
| 1 | 2 | 0.032 | 0.03 | 0.025 | 0.032329 | 0.0336 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033096 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.038133 | 0.032443 | #N/A
| 1 | 3 | 0.031 | 0.03 | 0.024 | 0.033519 | 0.031 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033841 | 0.03 | 0.03 | 0.03 | 0.03 | 0.04 | 0.03 | 0.02 | 0.034056 | 0.031741 | 16
| 1 | 4 | 0.032 | 0.03 | 0.023 | 0.030683 | 0.0323 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033586 | 0.02 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.039171 | 0.032099 | 0.030848
| 1 | 5 | 0.033 | 0.03 | 0.024 | 0.03098 | 0.0323 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033043 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.030944 | 0.033987 | 21
| 1 | 6 | 0.032 | 0.03 | 0.024 | 0.033836 | 0.031 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033272 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.039216 | 0.032259
| Mean |  | 0.032 | 0.028333 | 0.024 | 0.032144 | 0.032083 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033355 | 0.028333 | 0.03 | 0.03 | 0.03 | 0.031667 | 0.03 | 0.02 | 0.035068 | 0.032691
| Median |  | 0.032 | 0.03 | 0.024 | 0.031924 | 0.0323 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.033283 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02 | 0.036094 | 0.032351
| Std Dev. |  | 0.00063246 | 0.0040825 | 0.00063246 | 0.0013166 | 0.0009786 | 0 | 0 | 0 | 0 | 0 | 0.00030461 | 0.0040825 | 0 | 0 | 0 | 0.0040825 | 0 | 0 | 0.004464 | 0.00089924
| Rel.Std.Dev. |  | 1.976% | 14.41% | 2.635% | 4.096% | 3.05% | 0% | 0% | 0% | 0% | 0% | 0.9132% | 14.41% | 0% | 0% | 0% | 12.89% | 0% | 0% | 12.73% | 2.751%
| PDM3 |  | 3.735% | -8.151% | -22.2% | 4.203% | 4.005% | -2.749% | -2.749% | -2.749% | -2.749% | -2.749% | 8.129% | -8.151% | -2.749% | -2.749% | -2.749% | 2.654% | -2.749% | -35.17% | 13.68% | 5.976%
| Z-Score (Absolute) |  | 0.81 | 0.67 | 2.43 | 0.87 | 0.84 | 0 | 0 | 0 | 0 | 0 | 1.36 | 0.67 | 0 | 0 | 0 | 0.67 | 0 | 4.05 | 2.05 | 1.09
| Analytical results for Na2O in OREAS 45f (Certified Value 0.091 wt.%) | [21 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.103 | 0.09 | 0.07 | 0.05234 | 0.0998 | 0.09 | 0.07 | 0.1 | 0.08 | 0.09 | 0.07457 | 0.08 | 0.11 | 0.09 | 0.13 | 0.09 | < 0.1 | 0.09 | 0.073554 | 0.11515 | 1
| 1 | 2 | 0.083 | 0.08 | 0.08 | 0.051136 | 0.0944 | 0.1 | 0.08 | 0.11 | 0.08 | 0.09 | 0.080037 | 0.09 | 0.11 | 0.09 | 0.14 | 0.09 | < 0.1 | 0.09 | 0.070366 | 0.11477 | #N/A
| 1 | 3 | 0.104 | 0.09 | 0.09 | 0.04703 | 0.0863 | 0.09 | 0.09 | 0.1 | 0.09 | 0.1 | 0.079935 | 0.08 | 0.12 | 0.09 | 0.13 | 0.1 | < 0.1 | 0.09 | 0.071497 | 0.11779 | 16
| 1 | 4 | 0.108 | 0.1 | 0.08 | 0.0504 | 0.0863 | 0.09 | 0.09 | 0.11 | 0.09 | 0.1 | 0.081114 | 0.08 | 0.11 | 0.09 | 0.13 | 0.09 | < 0.1 | 0.09 | 0.070381 | 0.10934 | 0.091243
| 1 | 5 | 0.089 | 0.08 | 0.07 | 0.04818 | 0.093 | 0.09 | 0.07 | 0.09 | 0.08 | 0.09 | 0.078831 | 0.09 | 0.11 | 0.09 | 0.13 | 0.1 | < 0.1 | 0.09 | 0.060068 | 0.12021 | 22
| 1 | 6 | 0.108 | 0.1 | 0.08 | 0.04606 | 0.0971 | 0.09 | 0.09 | 0.12 | 0.08 | 0.09 | 0.080223 | 0.08 | 0.11 | 0.09 | 0.14 | 0.09 | < 0.1 | 0.09 | 0.068388 | 0.12266
| Mean |  | 0.099167 | 0.09 | 0.078333 | 0.049191 | 0.092817 | 0.091667 | 0.081667 | 0.105 | 0.083333 | 0.093333 | 0.079118 | 0.083333 | 0.11167 | 0.09 | 0.13333 | 0.093333 |  | 0.09 | 0.069042 | 0.11665
| Median |  | 0.1035 | 0.09 | 0.08 | 0.04929 | 0.0937 | 0.09 | 0.085 | 0.105 | 0.08 | 0.09 | 0.079986 | 0.08 | 0.11 | 0.09 | 0.13 | 0.09 |  | 0.09 | 0.070373 | 0.11647
| Std Dev. |  | 0.010572 | 0.0089443 | 0.0075277 | 0.0024761 | 0.0055604 | 0.0040825 | 0.0098319 | 0.010488 | 0.005164 | 0.005164 | 0.0023446 | 0.005164 | 0.0040825 | 1.5202e-17 | 0.005164 | 0.005164 |  | 1.5202e-17 | 0.0047094 | 0.004677
| Rel.Std.Dev. |  | 10.66% | 9.938% | 9.61% | 5.034% | 5.991% | 4.454% | 12.04% | 9.989% | 6.197% | 5.533% | 2.963% | 6.197% | 3.656% | 1.689e-14% | 3.873% | 5.533% |  | 1.689e-14% | 6.821% | 4.009%
| PDM3 |  | 8.684% | -1.363% | -14.15% | -46.09% | 1.724% | 0.4639% | -10.5% | 15.08% | -8.669% | 2.291% | -13.29% | -8.669% | 22.38% | -1.363% | 46.13% | 2.291% |  | -1.363% | -24.33% | 27.85%
| Z-Score (Absolute) |  | 0.71 | 0 | 0.9 | 3.14 | 0.21 | 0.13 | 0.64 | 1.16 | 0.51 | 0.26 | 0.84 | 0.51 | 1.67 | 0 | 3.34 | 0.26 | 3.08 | 0 | 1.62 | 2.05
| Analytical results for Ni in OREAS 45f (Certified Value 293 ppm) | [21 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 01 | 10 | 12 | 19 | 33 | 39 | 40 | [13 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [13 empty] | 0
0
| 1 | 1 | 440 | 316.0707 | 270 | 270 | 240 | 302.5743 | 274.0315 | [13 empty] | 1
| 1 | 2 | 346 | 337.053 | 280 | 260 | 240 | 294.0616 | 273.801 | [13 empty] | #N/A
| 1 | 3 | 283 | 395.442 | 270 | 350 | 240 | 286.2783 | 272.5468 | [13 empty] | 16
| 1 | 4 | 299 | 308.3694 | 280 | 270 | 240 | 302.2925 | 267.953 | [13 empty] | 292.6923
| 1 | 5 | 472 | 286.3654 | 270 | 310 | 310 | 310.6555 | 283.8389 | [13 empty] | 23
| 1 | 6 | 314 | 298.0747 | 270 | 290 | 310 | 286.2783 | 275.0173
| Mean |  | 359 | 323.5625 | 273.3333 | 291.6667 | 263.3333 | 297.0234 | 274.5314
| Median |  | 330 | 312.22 | 270 | 280 | 240 | 298.1771 | 273.9162
| Std Dev. |  | 78.613 | 39.1634 | 5.164 | 33.7145 | 36.1478 | 9.8397 | 5.1935
| Rel.Std.Dev. |  | 21.9% | 12.1% | 1.889% | 11.56% | 13.73% | 3.313% | 1.892%
| PDM3 |  | 22.65% | 10.55% | -6.614% | -0.3504% | -10.03% | 1.48% | -6.205%
| Z-Score (Absolute) |  | 2.13 | 0.97 | 0.69 | 0.09 | NA | 0.09 | 0.65
Results from laboratory 33 were removed due to their 100 ppm reading resolution.
| Analytical results for P2O5 in OREAS 45f (Certified Value 0.071 wt.%) | [21 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.068 | 0.07 | 0.074 | 0.078396 | 0.0756 | 0.07 | 0.05 | 0.0733 | 0.07 | 0.07 | 0.070791 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.063 | 0.07 | 0.062194 | 0.069755 | 1
| 1 | 2 | 0.073 | 0.07 | 0.074 | 0.074798 | 0.0733 | 0.07 | 0.05 | 0.0733 | 0.07 | 0.071 | 0.072269 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.065 | 0.07 | 0.066265 | 0.070818 | #N/A
| 1 | 3 | 0.072 | 0.07 | 0.075 | 0.073629 | 0.0756 | 0.07 | 0.06 | 0.071 | 0.07 | 0.071 | 0.072191 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.065 | 0.07 | 0.063243 | 0.070045 | 16
| 1 | 4 | 0.068 | 0.07 | 0.075 | 0.077456 | 0.071 | 0.07 | 0.06 | 0.0733 | 0.07 | 0.068 | 0.071649 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.063 | 0.07 | 0.063171 | 0.070043 | 0.070773
| 1 | 5 | 0.071 | 0.07 | 0.075 | 0.075714 | 0.0733 | 0.07 | 0.05 | 0.071 | 0.07 | 0.07 | 0.073184 | 0.07 | 0.07 | 0.07 | 0.07 | 0.07 | 0.062 | 0.07 | 0.062175 | 0.069143 | 24
| 1 | 6 | 0.069 | 0.07 | 0.074 | 0.076608 | 0.0733 | 0.07 | 0.06 | 0.0756 | 0.07 | 0.07 | 0.073387 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.066 | 0.07 | 0.067391 | 0.070089
| Mean |  | 0.070167 | 0.07 | 0.0745 | 0.0761 | 0.073683 | 0.07 | 0.055 | 0.072917 | 0.07 | 0.07 | 0.072245 | 0.073333 | 0.07 | 0.07 | 0.076667 | 0.07 | 0.064 | 0.07 | 0.064073 | 0.069982
| Median |  | 0.07 | 0.07 | 0.0745 | 0.076161 | 0.0733 | 0.07 | 0.055 | 0.0733 | 0.07 | 0.07 | 0.07223 | 0.07 | 0.07 | 0.07 | 0.08 | 0.07 | 0.064 | 0.07 | 0.063207 | 0.070044
| Std Dev. |  | 0.002137 | 0 | 0.00054772 | 0.0017503 | 0.0017314 | 0 | 0.0054772 | 0.0017314 | 0 | 0.0010954 | 0.00096562 | 0.005164 | 0 | 0 | 0.005164 | 0 | 0.0015492 | 0 | 0.0022111 | 0.0005428
| Rel.Std.Dev. |  | 3.046% | 0% | 0.7352% | 2.3% | 2.35% | 0% | 9.959% | 2.374% | 0% | 1.565% | 1.337% | 7.042% | 0% | 0% | 6.736% | 0% | 2.421% | 0% | 3.451% | 0.7756%
| PDM3 |  | -0.8567% | -1.092% | 5.266% | 7.527% | 4.112% | -1.092% | -22.29% | 3.029% | -1.092% | -1.092% | 2.08% | 3.618% | -1.092% | -1.092% | 8.328% | -1.092% | -9.57% | -1.092% | -9.467% | -1.117%
| Z-Score (Absolute) |  | 0.33 | NA | 1.05 | 1.55 | 0.8 | NA | NA | 0.55 | NA | 0.38 | 0.33 | NA | NA | NA | NA | NA | 2.29 | NA | 2.26 | 0.39
Results from laboratories 5, 14, 17, 20, 25, 28, 29, 31, 33 & 35 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Rb in OREAS 45f (Indicative Value 36.9 ppm) | [21 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | [18 empty] | 1
| No | No | 33 | 39 | [18 empty] | ppm
|  |  | BF*XRF | BF*XRF | [18 empty] | 1
1
| 1 | 1 | < 90 | 28.4021 | [18 empty] | 1
| 1 | 2 | < 90 | 37.8263 | [18 empty] | 2
| 1 | 3 | < 90 | 37.8773 | [18 empty] | 16
| 1 | 4 | < 90 | 37.8342 | [18 empty] | 36.8733
| 1 | 5 | < 90 | 37.8577 | [18 empty] | 8
| 1 | 6 | < 90 | 41.4422
| Mean |  |  | 36.8733
| Median |  |  | 37.8459
| Std Dev. |  |  | 4.3919
| Rel.Std.Dev. |  |  | 11.91%
| PDM3 |  |  | -2.22e-14%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 45f (Certified Value 0.032 wt.%) | [21 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 07 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 28 | 29 | 31 | 35 | 39 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.0316 | 0.03 | 0.034 | 0.032 | < 0.004 | 0.03 | 0.0288 | 0.028 | 0.031031 | 0.024 | 0.028 | 0.032 | 0.032 | 0.023617 | [6 empty] | 1
| 1 | 2 | 0.032 | 0.03 | 0.036 | 0.032 | < 0.004 | 0.032 | 0.03 | 0.032 | 0.031222 | 0.024 | 0.028 | 0.032 | 0.032 | 0.021107 | [6 empty] | 28
| 1 | 3 | 0.0336 | 0.03 | 0.033 | 0.028 | < 0.004 | 0.034 | 0.0292 | 0.032 | 0.031037 | 0.024 | 0.028 | 0.032 | 0.032 | 0.024451 | [6 empty] | 16
| 1 | 4 | 0.032 | 0.03 | 0.035 | 0.028 | < 0.004 | 0.032 | 0.03 | 0.032 | 0.030231 | 0.028 | 0.028 | 0.032 | 0.032 | 0.022353 | [6 empty] | 0.031501
| 1 | 5 | 0.0316 | 0.03 | 0.034 | 0.028 | < 0.004 | 0.029 | 0.0296 | 0.032 | 0.032328 | 0.024 | 0.028 | 0.032 | 0.032 | 0.023196 | [6 empty] | 25
| 1 | 6 | 0.032 | 0.032 | 0.033 | 0.032 | < 0.004 | 0.033 | 0.0292 | 0.032 | 0.033754 | 0.024 | 0.028 | 0.032 | 0.032 | 0.025694
| Mean |  | 0.032133 | 0.030333 | 0.034167 | 0.03 |  | 0.031667 | 0.029467 | 0.031333 | 0.0316 | 0.024667 | 0.028 | 0.032 | 0.032 | 0.023403
| Median |  | 0.032 | 0.03 | 0.034 | 0.03 |  | 0.032 | 0.0294 | 0.032 | 0.03113 | 0.024 | 0.028 | 0.032 | 0.032 | 0.023406
| Std Dev. |  | 0.00074476 | 0.0008165 | 0.001169 | 0.0021909 |  | 0.0018619 | 0.00048442 | 0.001633 | 0.0012511 | 0.001633 | 0 | 0 | 0 | 0.0016006
| Rel.Std.Dev. |  | 2.318% | 2.692% | 3.422% | 7.303% |  | 5.88% | 1.644% | 5.212% | 3.959% | 6.62% | 0% | 0% | 0% | 6.839%
| PDM3 |  | 2.008% | -3.706% | 8.462% | -4.765% |  | 0.5262% | -6.458% | -0.532% | 0.3158% | -21.7% | -11.11% | 1.584% | 1.584% | -25.71%
| Z-Score (Absolute) |  | 0.61 | 0.32 | 1.62 | NA | 14.63 | 0.35 | 0.74 | NA | 0.32 | NA | NA | NA | NA | 3.82
Results from laboratories 14, 23, 28, 29, 31 & 35 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for SiO2 in OREAS 45f (Certified Value 43.59 wt.%) | [21 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 43.613 | 44.2 | 43.59 | 43.5019 | 43.69 | 43.53 | 42.31 | 43.71 | 43.4 | 43.8 | 44.8133 | 43.4 | 43.79 | 43.51 | 43.23 | 43.7 | 43.3 | 43.18 | 42.3914 | 42.455 | 1
| 1 | 2 | 43.741 | 44 | 43.63 | 43.4774 | 43.72 | 43.62 | 42.35 | 43.8 | 43.5 | 43.97 | 45.246 | 43.2 | 43.78 | 43.66 | 43.71 | 43.8 | 43.6 | 43.31 | 42.2263 | 42.433 | #N/A
| 1 | 3 | 43.602 | 44 | 43.42 | 43.3282 | 43.82 | 43.58 | 42.28 | 43.79 | 43.5 | 43.92 | 45.3281 | 43.3 | 43.83 | 43.75 | 43.36 | 43.7 | 42.8 | 43.22 | 42.2697 | 42.7181 | 16
| 1 | 4 | 43.814 | 44 | 43.41 | 43.4526 | 43.83 | 43.55 | 42.3 | 43.69 | 43.5 | 43.78 | 45.3691 | 43.2 | 44.08 | 43.37 | 43.39 | 43.7 | 43 | 43.4 | 41.5947 | 42.6053 | 43.5888
| 1 | 5 | 43.51 | 44.2 | 43.54 | 43.2966 | 43.72 | 43.62 | 42.24 | 43.7 | 43.6 | 43.86 | 45.2767 | 43.4 | 43.46 | 43.58 | 43.53 | 43.7 | 43.5 | 43.33 | 41.9311 | 42.7834 | 26
| 1 | 6 | 43.791 | 44 | 43.62 | 43.3586 | 43.66 | 43.7 | 42.43 | 43.66 | 43.4 | 43.87 | 45.2637 | 43.2 | 43.24 | 43.77 | 43.52 | 43.8 | 42.7 | 43.4 | 41.7001 | 42.2416
| Mean |  | 43.6785 | 44.0667 | 43.535 | 43.4025 | 43.74 | 43.6 | 42.3183 | 43.725 | 43.4833 | 43.8667 | 45.2162 | 43.2833 | 43.6967 | 43.6067 | 43.4567 | 43.7333 | 43.15 | 43.3067 | 42.0189 | 42.5394
| Median |  | 43.677 | 44 | 43.565 | 43.4056 | 43.72 | 43.6 | 42.305 | 43.705 | 43.5 | 43.865 | 45.2702 | 43.25 | 43.785 | 43.62 | 43.455 | 43.7 | 43.15 | 43.32 | 42.0787 | 42.5301
| Std Dev. |  | 0.12121 | 0.10328 | 0.098133 | 0.085616 | 0.06957 | 0.060992 | 0.065549 | 0.056833 | 0.075277 | 0.071461 | 0.20248 | 0.098319 | 0.29831 | 0.1524 | 0.16657 | 0.05164 | 0.37283 | 0.091141 | 0.3268 | 0.20144
| Rel.Std.Dev. |  | 0.2775% | 0.2344% | 0.2254% | 0.1973% | 0.1591% | 0.1399% | 0.1549% | 0.13% | 0.1731% | 0.1629% | 0.4478% | 0.2272% | 0.6827% | 0.3495% | 0.3833% | 0.1181% | 0.864% | 0.2105% | 0.7777% | 0.4735%
| PDM3 |  | 0.2058% | 1.096% | -0.1235% | -0.4273% | 0.3468% | 0.02566% | -2.915% | 0.3124% | -0.242% | 0.6374% | 3.733% | -0.7008% | 0.2474% | 0.04095% | -0.3032% | 0.3315% | -1.007% | -0.6473% | -3.602% | -2.408%
| Z-Score (Absolute) |  | 0.44 | 1.99 | 0.13 | 0.66 | 0.69 | 0.13 | 4.98 | 0.63 | 0.34 | 1.19 | 6.57 | 1.13 | 0.51 | 0.16 | 0.44 | 0.66 | 1.66 | 1.04 | 6.17 | 4.1
| Analytical results for Sn in OREAS 45f (Indicative Value < 10 ppm) | [21 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 01 | 12 | 19 | 33 | 39 | [15 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [15 empty] | 1
1
| 1 | 1 | 8 | < 10 | < 10 | < 80 | < 31.08 | [15 empty] | 1
| 1 | 2 | < 8 | < 10 | < 10 | < 80 | < 31.04 | [15 empty] | 3
| 1 | 3 | < 8 | < 10 | < 10 | < 80 | < 31.08 | [15 empty] | 16
| 1 | 4 | < 8 | < 10 | < 10 | < 80 | < 31.05 | [15 empty] | < 10
| 1 | 5 | 39 | < 10 | < 10 | < 80 | < 31.07 | [15 empty] | 9
| 1 | 6 | 24 | < 10 | < 10 | < 80 | < 31.08
| Mean |  | 23.6667
| Median |  | 24
| Std Dev. |  | 15.5027
| Rel.Std.Dev. |  | 65.5%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 0.67 | 1.97 | 0.13
| Analytical results for Sr in OREAS 45f (Certified Value 31.4 ppm) | [21 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 10 | 12 | 14 | 17 | 19 | 20 | 28 | 29 | 31 | 33 | 35 | 39 | 40 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 1
1
| 1 | 1 | 59 | 35.09 | 40 | < 80 | < 80 | 40 | 30 | 80 | < 80 | < 80 | < 80 | 80 | 20.7168 | 27.5335 | [6 empty] | 1
| 1 | 2 | 68 | 34.42 | 20 | < 80 | < 80 | 50 | 30 | 80 | < 80 | < 80 | < 80 | 80 | 41.3864 | 25.8469 | [6 empty] | #N/A
| 1 | 3 | < 8 | 29.51 | 40 | < 80 | < 80 | 40 | 30 | 80 | < 80 | < 80 | < 80 | 80 | 41.4422 | 25.0688 | [6 empty] | 16
| 1 | 4 | < 8 | 38.22 | 30 | < 80 | < 80 | 60 | 30 | 80 | < 80 | < 80 | < 80 | 80 | 31.0463 | 27.6486 | [6 empty] | 31.3732
| 1 | 5 | < 8 | 40.33 | 40 | < 80 | < 80 | 40 | 30 | 80 | < 80 | < 80 | < 80 | 80 | 31.0655 | 25.6973 | [6 empty] | 27
| 1 | 6 | < 8 | 29.34 | 20 | < 80 | < 80 | 30 | 30 | 80 | < 80 | < 80 | < 80 | 80 | 41.4422 | 25.3847
| Mean |  | 63.5 | 34.485 | 31.6667 |  |  | 43.3333 | 30 | 80 |  |  |  | 80 | 34.5166 | 26.1966
| Median |  | 63.5 | 34.755 | 35 |  |  | 40 | 30 | 80 |  |  |  | 80 | 36.226 | 25.7721
| Std Dev. |  | 6.364 | 4.4656 | 9.8319 |  |  | 10.328 | 0 | 0 |  |  |  | 0 | 8.4559 | 1.1134
| Rel.Std.Dev. |  | 10.02% | 12.95% | 31.05% |  |  | 23.83% | 0% | 0% |  |  |  | 0% | 24.5% | 4.25%
| PDM3 |  | 102.4% | 9.919% | 0.9354% |  |  | 38.12% | -4.377% | 155% |  |  |  | 155% | 10.02% | -16.5%
| Z-Score (Absolute) |  | NA | 4.99 | 6.79 | 0 | 0 | 0.67 | 7.86 | NA | 0 | 0 | 0 | NA | 4.97 | 10.29
Results from laboratory 1 were removed due to their lack of accuracy Results from laboratories 28 & 35 were removed due to their 100 ppm reading resolution.
| Analytical results for TiO2 in OREAS 45f (Certified Value 2 wt.%) | [21 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 39 | 40 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.051 | 2.02 | 2 | 2.0177 | 2 | 1.98 | 2.2 | 1.99 | 1.96 | 1.99 | 2.0034 | 1.98 | 1.99 | 2.01 | 1.97 | 2 | 1.97 | 1.98 | 2.0272 | 1.9789 | 1
| 1 | 2 | 2.055 | 2.03 | 2 | 2.0192 | 1.98 | 1.98 | 2.19 | 1.99 | 1.96 | 2.01 | 2.0316 | 1.97 | 2.02 | 2.01 | 2.01 | 1.99 | 2 | 1.98 | 2.0167 | 1.9789 | #N/A
| 1 | 3 | 2.064 | 2.04 | 1.99 | 2.0269 | 2 | 1.98 | 2.18 | 1.98 | 1.97 | 2 | 2.02 | 1.99 | 2 | 2.01 | 1.98 | 2 | 1.98 | 1.98 | 2.0399 | 1.9859 | 16
| 1 | 4 | 2.051 | 2.04 | 1.99 | 2.0125 | 2 | 1.98 | 2.2 | 1.98 | 1.96 | 2 | 2.0242 | 1.99 | 2.02 | 2 | 1.98 | 1.98 | 1.99 | 1.97 | 1.9852 | 1.9853 | 1.9952
| 1 | 5 | 2.055 | 2.04 | 2 | 1.9916 | 1.98 | 1.99 | 2.18 | 1.99 | 1.97 | 1.99 | 2.0322 | 1.98 | 1.97 | 2.01 | 1.99 | 2 | 2 | 1.98 | 1.9906 | 1.9927 | 28
| 1 | 6 | 2.049 | 2.03 | 2 | 2.0217 | 1.98 | 1.99 | 2.19 | 1.99 | 1.97 | 2.01 | 2.0172 | 1.98 | 1.96 | 2.02 | 2 | 1.99 | 1.96 | 1.99 | 1.9968 | 1.9747
| Mean |  | 2.0542 | 2.0333 | 1.9967 | 2.0149 | 1.99 | 1.9833 | 2.19 | 1.9867 | 1.965 | 2 | 2.0214 | 1.9817 | 1.9933 | 2.01 | 1.9883 | 1.9933 | 1.9833 | 1.98 | 2.0094 | 1.9827
| Median |  | 2.053 | 2.035 | 2 | 2.0185 | 1.99 | 1.98 | 2.19 | 1.99 | 1.965 | 2 | 2.0221 | 1.98 | 1.995 | 2.01 | 1.985 | 1.995 | 1.985 | 1.98 | 2.0067 | 1.9821
| Std Dev. |  | 0.0053821 | 0.008165 | 0.005164 | 0.012382 | 0.010954 | 0.005164 | 0.0089443 | 0.005164 | 0.0054772 | 0.0089443 | 0.010696 | 0.0075277 | 0.025033 | 0.0063246 | 0.01472 | 0.008165 | 0.01633 | 0.0063246 | 0.021905 | 0.0064656
| Rel.Std.Dev. |  | 0.262% | 0.4016% | 0.2586% | 0.6145% | 0.5505% | 0.2604% | 0.4084% | 0.2599% | 0.2787% | 0.4472% | 0.5291% | 0.3799% | 1.256% | 0.3147% | 0.7403% | 0.4096% | 0.8234% | 0.3194% | 1.09% | 0.3261%
| PDM3 |  | 2.956% | 1.912% | 0.07382% | 0.9896% | -0.2603% | -0.5944% | 9.764% | -0.4274% | -1.513% | 0.2409% | 1.315% | -0.678% | -0.09324% | 0.7421% | -0.3438% | -0.09324% | -0.5944% | -0.7615% | 0.7117% | -0.6247%
| Z-Score (Absolute) |  | 3.68 | 2.42 | 0.2 | 1.31 | 0.2 | 0.61 | 11.91 | 0.4 | 1.72 | 0.4 | 1.7 | 0.71 | 0 | 1.01 | 0.3 | 0 | 0.61 | 0.81 | 0.97 | 0.64
| Analytical results for V2O5 in OREAS 45f (Certified Value 447 ppm) | [21 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 05 | 10 | 12 | 19 | 24 | 25 | 33 | 39 | 40 | [10 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [10 empty] | 0
0
| 1 | 1 | 520 | 600 | 410.8 | 518 | 411 | 364.9405 | 500 | 600 | 536.2627 | 466.8632 | [10 empty] | 1
| 1 | 2 | 480 | 500 | 412.4 | 482 | 428 | 354.0894 | 500 | 600 | 591.0647 | 464.2746 | [10 empty] | #N/A
| 1 | 3 | 560 | 400 | 409.1 | 482 | 411 | 360.0801 | 500 | 600 | 459.9787 | 473.1954 | [10 empty] | 16
| 1 | 4 | 520 | 500 | 424.1 | 518 | 411 | 370.119 | 500 | 500 | 584.7611 | 462.3785 | [10 empty] | 447.2114
| 1 | 5 | 540 | 500 | 419.8 | 482 | 411 | 366.3675 | 400 | 600 | 499.124 | 471.8778 | [10 empty] | 29
| 1 | 6 | 530 | 500 | 443.1 | 482 | 446 | 396.5655 | 500 | 600 | 512.2491 | 464.5728
| Mean |  | 525 | 500 | 419.8833 | 494 | 419.6667 | 368.6937 | 483.3333 | 583.3333 | 530.5734 | 467.1937
| Median |  | 525 | 500 | 416.1 | 482 | 411 | 365.654 | 500 | 600 | 524.2559 | 465.718
| Std Dev. |  | 26.6458 | 63.2456 | 12.7404 | 18.5903 | 14.5831 | 14.7367 | 40.8248 | 40.8248 | 50.8652 | 4.3965
| Rel.Std.Dev. |  | 5.075% | 12.65% | 3.034% | 3.763% | 3.475% | 3.997% | 8.447% | 6.999% | 9.587% | 0.941%
| PDM3 |  | 17.39% | 11.8% | -6.111% | 10.46% | -6.159% | -17.56% | 8.077% | 30.44% | 18.64% | 4.468%
| Z-Score (Absolute) |  | 1.48 | NA | 0.43 | 0.91 | 0.44 | 1.36 | NA | NA | NA | 0.43
Results from laboratories 5, 25, 33 & 39 were removed due to their 100 ppm reading resolution.
| Analytical results for Zn in OREAS 45f (Indicative Value 26.4 ppm) | [21 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 01 | 10 | 12 | 33 | 39 | 40 | [14 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [14 empty] | 1
1
| 1 | 1 | 24 | 32.696 | 30 | < 80 | 41.4336 | 51.5343 | [14 empty] | 1
| 1 | 2 | 24 | 14.2995 | 20 | < 80 | 41.3864 | 49.9773 | [14 empty] | 1
| 1 | 3 | 24 | 27.876 | 30 | < 80 | 41.4422 | 51.1357 | [14 empty] | 16
| 1 | 4 | 24 | 31.1697 | 30 | < 80 | 41.395 | 47.5286 | [14 empty] | 26.3593
| 1 | 5 | 24 | 25.7873 | 30 | < 80 | 31.0655 | 50.4864 | [14 empty] | 7
| 1 | 6 | 32 | 20.0032 | 30 | < 80 | 41.4422 | 50.0146
| Mean |  | 25.3333 | 25.3053 | 28.3333 |  | 39.6942 | 50.1128
| Median |  | 24 | 26.8316 | 30 |  | 41.4143 | 50.2505
| Std Dev. |  | 3.266 | 7.0013 | 4.0825 |  | 4.2272 | 1.408
| Rel.Std.Dev. |  | 12.89% | 27.67% | 14.41% |  | 10.65% | 2.81%
| PDM3 |  | -3.892% | -3.999% | 7.489% |  | 50.59% | 90.11%
| Z-Score (Absolute) |  | 0.79 | 0.8 | 0.52 | 0.56 | 0.52 | 1.47
| Analytical results for Zr in OREAS 45f (Certified Value 263 ppm) | [21 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 10 | 12 | 19 | 24 | 33 | 39 | 40 | [12 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [12 empty] | 0
0
| 1 | 1 | 274 | 266.9529 | 250 | 290 | 188.4126 | 300 | 244.5262 | 264.1345 | [12 empty] | 1
| 1 | 2 | 267 | 250.4442 | 240 | 310 | 192.1674 | 220 | 251.8806 | 262.0728 | [12 empty] | #N/A
| 1 | 3 | 281 | 254.6639 | 250 | 280 | 194.5802 | 300 | 259.8629 | 265.4434 | [12 empty] | 16
| 1 | 4 | 289 | 255.3302 | 250 | 290 | 195.7682 | 220 | 259.5671 | 260.7149 | [12 empty] | 263.0773
| 1 | 5 | 267 | 251.1105 | 250 | 290 | 194.5201 | 300 | 252.0893 | 266.4137 | [12 empty] | 30
| 1 | 6 | 281 | 250.2221 | 240 | 280 | 196.2239 | 220 | 229.2908 | 262.3833
| Mean |  | 276.5 | 254.7873 | 246.6667 | 290 | 193.6121 | 260 | 249.5361 | 263.5271
| Median |  | 277.5 | 252.8872 | 250 | 290 | 194.5502 | 260 | 251.985 | 263.2589
| Std Dev. |  | 8.7579 | 6.3488 | 5.164 | 10.9545 | 2.9105 | 43.8178 | 11.439 | 2.1774
| Rel.Std.Dev. |  | 3.167% | 2.492% | 2.094% | 3.777% | 1.503% | 16.85% | 4.584% | 0.8262%
| PDM3 |  | 5.102% | -3.151% | -6.238% | 10.23% | -26.4% | -1.17% | -5.147% | 0.171%
| Z-Score (Absolute) |  | 1.67 | 0 | 0.63 | 2.72 | 4.72 | NA | 0.41 | 0.67
Results from laboratory 33 were removed due to their 100 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 45f (Indicative Value 3.95 wt.%) | [28 empty] | Indicative
H2O- | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 10 | 12 | 14 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 39 | 40 | wt.%
|  |  | LOI*TGA | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | LOI*TGA | at 105°C*Oven | at 105°C*TGA | at 105°C*Oven | LOI*TGA | LOI*TGA | at 105°C*Oven | LOI*TGA | at 105°C*Oven | at 105°C*Oven | LOI*TGA | at 105°C*Oven | at 105°C*Oven | LOI*TGA | BF*XRF | at 105°C*TGA | at 105°C*TGA | at 105°C*Oven | at 105°C*Oven | at 105°C*Oven | BF*XRF | 2
2
| 1 | 1 | 4.57 | 5.407 | 4.4 | 3.71 | 3.6 | 2.6652 | 3.36 | 4.02 | 2.07 | 3.17 | 4.09 | 4.1 | 3.81 | 3.11 | 4.51 | 3.47 | 4.86 | 3.93 | 2.97 | 4.54 | 5.25 | 4.61 | 3.53 | 4.11 | 4.18 | 3.46 | 3.4289 | 1
| 1 | 2 | 4.97 | 5.2773 | 4.5 | 3.69 | 3.7 | 2.7891 | 3.3 | 3.95 | 2.05 | 3.29 | 4.08 | 4.15 | 3.84 | 3.21 | 4.51 | 3.14 | 4.85 | 3.63 | 3.07 | 4.5 | 5.14 | 4.69 | 3.4 | 4.14 | 4.17 | 3.35 | 3.3311 | 5
| 1 | 3 | 4.64 | 5.3545 | 4.5 | 3.62 | 3.8 | 2.6162 | 3.38 | 4.1 | 2.1 | 3.15 | 4.05 | 4.04 | 3.82 | 3.47 | 4.7 | 3.38 | 4.87 | 3.82 | 3.11 | 4.32 | 5.14 | 4.61 | 3.49 | 4.06 | 4.18 | 3.48 | 3.4487 | 16
| 1 | 4 | 4.52 | 5.2868 | 4.5 | 3.66 | 3.9 | 2.7903 | 3.35 | 4.01 | 2.08 | 3.34 | 4.1 | 4.09 | 3.74 | 3.29 | 4.72 | 3.41 | 4.79 | 3.34 | 3.02 | 4.54 | 5.13 | 4.56 | 3.49 | 3.89 | 4.21 | 3.37 | 3.3492 | 3.9513
| 1 | 5 | 4.72 | 5.4057 | 4.6 | 3.66 | 3.9 | 2.5726 | 3.35 | 3.96 | 2.06 | 3.33 | 4.09 | 4 | 3.78 | 3.44 | 4.35 | 3.44 | 4.79 | 3.72 | 3.02 | 4.1 | 5.1 | 4.61 | 3.45 | 4.11 | 4.1 | 3.43 | 3.4254 | 11
| 1 | 6 | 4.99 | 5.2921 | 4.5 | 3.67 | 3.8 | 2.6052 | 3.38 | 3.96 | 2.1 | 3.17 | 4.09 | 3.95 | 3.74 | 3.31 | 4.48 | 3.44 | 4.84 | 3.93 | 3.03 | 4.15 | 5.28 | 4.72 | 3.43 | 3.87 | 4.1 | 3.48 | 3.3858
| Mean |  | 4.735 | 5.3372 | 4.5 | 3.6683 | 3.7833 | 2.6731 | 3.3533 | 4 | 2.0767 | 3.2417 | 4.0833 | 4.055 | 3.7883 | 3.305 | 4.545 | 3.38 | 4.8333 | 3.7283 | 3.0367 | 4.3583 | 5.1733 | 4.6333 | 3.465 | 4.03 | 4.1567 | 3.4283 | 3.3948
| Median |  | 4.68 | 5.3233 | 4.5 | 3.665 | 3.8 | 2.6407 | 3.355 | 3.985 | 2.075 | 3.23 | 4.09 | 4.065 | 3.795 | 3.3 | 4.51 | 3.425 | 4.845 | 3.77 | 3.025 | 4.41 | 5.14 | 4.61 | 3.47 | 4.085 | 4.175 | 3.445 | 3.4056
| Std Dev. |  | 0.20147 | 0.060049 | 0.063246 | 0.030605 | 0.1169 | 0.095085 | 0.029439 | 0.056921 | 0.020656 | 0.087731 | 0.017512 | 0.07287 | 0.042151 | 0.1362 | 0.14096 | 0.12149 | 0.035024 | 0.22373 | 0.048028 | 0.19884 | 0.073121 | 0.059554 | 0.047223 | 0.11916 | 0.045898 | 0.056362 | 0.047376
| Rel.Std.Dev. |  | 4.255% | 1.125% | 1.405% | 0.8343% | 3.09% | 3.557% | 0.8779% | 1.423% | 0.9947% | 2.706% | 0.4289% | 1.797% | 1.113% | 4.121% | 3.101% | 3.594% | 0.7246% | 6.001% | 1.582% | 4.562% | 1.413% | 1.285% | 1.363% | 2.957% | 1.104% | 1.644% | 1.396%
| PDM3 |  | 19.83% | 35.07% | 13.89% | -7.162% | -4.252% | -32.35% | -15.13% | 1.231% | -47.44% | -17.96% | 3.34% | 2.623% | -4.125% | -16.36% | 15.02% | -14.46% | 22.32% | -5.644% | -23.15% | 10.3% | 30.93% | 17.26% | -12.31% | 1.991% | 5.196% | -13.24% | -14.08%
| Z-Score (Absolute) |  | 1.47 | 2.4 | 1.1 | 0.19 | 0.01 | 1.73 | 0.67 | 0.33 | 2.65 | 0.85 | 0.46 | 0.41 | 0 | 0.75 | 1.17 | 0.63 | 1.62 | 0.09 | 1.17 | 0.88 | 2.15 | 1.31 | 0.5 | 0.37 | 0.57 | 0.56 | 0.61
| Analytical results for LOI1000 in OREAS 45f (Certified Value 10.13 wt.%) | [28 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 10 | 12 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 31 | 33 | 34 | 35 | 36 | 39 | 40 |  |  |  | wt.%
|  |  | LOI*TGA | at 105°C*Oven | LOI*TGA | BF*XRF | BF*XRF | LOI*TGA | LOI*TGA | LOI*Furnace | BF*XRF | LOI*TGA | BF*XRF | BF*OES | LOI*TGA | LOI*Furnace | BF*XRF | LOI*Furnace | LOI*Furnace | LOI*TGA | BF*XRF | LOI*Furnace | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF |  |  |  | 2
3
| 1 | 1 | 10.2 | 9.9006 | 10.36 | 10.3 | 10.46 | 10.2023 | 10.23 | 10.12 | 11.02 | 9.95 | 10 | 10.39 | 9.99 | 9.9324 | 11 | 10.32 | 10.38 | 10.09 | 10.1 | 10.19 | 10.23 | 9.5839 | 9.7162 | 10.1094 |  |  |  | 1
| 1 | 2 | 10.15 | 9.9515 | 10.16 | 10.3 | 10.45 | 10.571 | 10.17 | 10.06 | 10.96 | 9.95 | 9.9 | 10.15 | 9.95 | 9.8766 | 11.1 | 10.02 | 10.37 | 10.1 | 10.1 | 10.23 | 10.24 | 9.6599 | 9.7672 | 10.129 |  |  |  | #N/A
| 1 | 3 | 10.12 | 9.8753 | 10.47 | 10.4 | 10.45 | 10.3415 | 10.19 | 10.02 | 10.89 | 9.86 | 10 | 10.41 | 10.03 | 9.8859 | 11.1 | 10.31 | 10.39 | 10.12 | 10.1 | 10.29 | 10.16 | 9.631 | 9.6975 | 10.0832 |  |  |  | 16
| 1 | 4 | 10.12 | 10.12 | 10.26 | 10.2 | 10.48 | 10.51 | 10.17 | 9.99 | 10.89 | 9.96 | 10 | 10.53 | 10.03 | 9.9138 | 11.2 | 10.3 | 10.44 | 10.1 | 10.1 | 10.11 | 10.14 | 9.7909 | 9.8106 | 10.0859 |  |  |  | 10.1255
| 1 | 5 | 10.11 | 9.8658 | 10.27 | 10.3 | 10.47 | 9.979 | 10.17 | 10.03 | 10.99 | 9.93 | 10 | 10.52 | 10 | 9.9231 | 11.2 | 10.23 | 10.32 | 10.13 | 10.1 | 10.55 | 10.18 | 9.6882 | 9.7856 | 10.0865 |  |  |  | 32
| 1 | 6 | 10.17 | 9.7056 | 10.37 | 9.92 | 10.5 | 10.897 | 10.21 | 10.06 | 10.89 | 9.92 | 10.1 | 10.22 | 10.03 | 9.8766 | 11.1 | 10.29 | 10.35 | 10.11 | 10.1 | 10.3 | 10.2 | 9.9137 | 9.7286 | 10.1427
| Mean |  | 10.145 | 9.9031 | 10.315 | 10.2367 | 10.4683 | 10.4168 | 10.19 | 10.0467 | 10.94 | 9.9283 | 10 | 10.37 | 10.005 | 9.9014 | 11.1167 | 10.245 | 10.375 | 10.1083 | 10.1 | 10.2783 | 10.1917 | 9.7113 | 9.7509 | 10.1061
| Median |  | 10.135 | 9.888 | 10.315 | 10.3 | 10.465 | 10.4257 | 10.18 | 10.045 | 10.925 | 9.94 | 10 | 10.4 | 10.015 | 9.8999 | 11.1 | 10.295 | 10.375 | 10.105 | 10.1 | 10.26 | 10.19 | 9.6741 | 9.7479 | 10.098
| Std Dev. |  | 0.035071 | 0.13459 | 0.10784 | 0.16753 | 0.019408 | 0.31824 | 0.025298 | 0.044572 | 0.057966 | 0.03656 | 0.063246 | 0.15556 | 0.032094 | 0.024722 | 0.075277 | 0.11467 | 0.040373 | 0.01472 | 0 | 0.15026 | 0.0392 | 0.12094 | 0.043815 | 0.025264
| Rel.Std.Dev. |  | 0.3457% | 1.359% | 1.045% | 1.637% | 0.1854% | 3.055% | 0.2483% | 0.4437% | 0.5298% | 0.3682% | 0.6325% | 1.5% | 0.3208% | 0.2497% | 0.6772% | 1.119% | 0.3891% | 0.1456% | 0% | 1.462% | 0.3846% | 1.245% | 0.4493% | 0.25%
| PDM3 |  | 0.1929% | -2.196% | 1.872% | 1.098% | 3.386% | 2.877% | 0.6373% | -0.7782% | 8.044% | -1.947% | -1.239% | 2.415% | -1.19% | -2.213% | 9.789% | 1.181% | 2.464% | -0.1692% | -0.2515% | 1.51% | 0.6538% | -4.091% | -3.699% | -0.1911%
| Z-Score (Absolute) |  | 0.09 | 1.08 | 0.6 | 0.28 | 1.23 | 1.02 | 0.09 | 0.49 | 3.16 | 0.98 | 0.68 | 0.83 | 0.66 | 1.09 | 3.88 | 0.32 | 0.85 | 0.24 | 0.28 | 0.45 | 0.1 | 1.86 | 1.7 | 0.25
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 45f (Indicative Value < 1 ppm) | [20 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 05 | 14 | 17 | 19 | 25 | 33 | 39 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | [12 empty] | 2
2
| 1 | 1 | < 1 | < 5 |  | < 5 | < 1 | < 10 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | < 5 |  | < 5 | < 1 | < 10 | < 1 | [12 empty] | 7
| 1 | 3 | < 1 | < 5 |  | < 5 | < 1 | < 10 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | < 5 |  | < 5 | < 1 | < 10 | < 1 | [12 empty] | < 1
| 1 | 5 | < 1 | 7 |  | < 5 | < 1 | < 10 | < 1 | [12 empty] | 13
| 1 | 6 | < 1 | 7 |  | < 5 | < 1 | < 10 | < 1
| Mean |  |  | 7
| Median |  |  | 7
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.67 | 1.69 | NA | 0.67 | 0.67 | 2.36 | 0.67
| Analytical results for Al in OREAS 45f (Certified Value 10.41 wt.%) | [20 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 10 | 13 | 15 | 17 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | BF*OES | BF*OES/MS | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  |  |  | 2
3
| 1 | 1 | 10.2047 | 10.553 | 10.6 | 9.8463 | 10.59 | 10.15 |  | 10.532 | 10.225 | 10.43 | 10.4066 | 10.4 | 10.4 | 10.744 | 10.1085 | 10.5727 |  |  |  | 1
| 1 | 2 | 10.1767 | 10.738 | 10.6 | 9.738 | 10.77 | 10.1522 |  | 10.505 | 10.273 | 10.15 | 10.2804 | 10.4 | 10.5 | 10.505 | 10.1085 | 10.1451 |  |  |  | #N/A
| 1 | 3 | 10.1883 | 10.352 | 10.5 | 9.7428 | 10.44 | 10.2533 |  | 10.574 | 10.357 | 9.96 | 10.0633 | 10.3 | 10.3 | 10.802 | 10.7436 | 10.4366 |  |  |  | 16
| 1 | 4 | 10.1418 | 10.654 | 10.5 | 9.8054 | 10.48 | 10.1256 |  | 10.638 | 10.283 | 10.46 | 10.0153 | 10.2 | 10.6 | 10.654 | 10.7965 | 10.4047 |  |  |  | 10.4102
| 1 | 5 | 10.1655 | 10.95 | 10.6 | 9.8487 | 10.89 | 10.2 |  | 10.484 | 10.352 | 10.19 | 10.501 | 10.4 | 10.3 | 10.643 | 10.2673 | 10.4567 |  |  |  | 34
| 1 | 6 | 10.2095 | 10.516 | 10.4 | 9.7765 | 10.52 | 10.3526 |  | 10.585 | 10.325 | 10.18 | 10.445 | 10.3 | 10.5 | 10.611 | 10.2673 | 9.9154
| Mean |  | 10.1811 | 10.6272 | 10.5333 | 9.793 | 10.615 | 10.2056 |  | 10.553 | 10.3025 | 10.2283 | 10.2853 | 10.3333 | 10.4333 | 10.6598 | 10.3819 | 10.3219
| Median |  | 10.1825 | 10.6035 | 10.55 | 9.7909 | 10.555 | 10.1761 |  | 10.553 | 10.304 | 10.185 | 10.3435 | 10.35 | 10.45 | 10.6485 | 10.2673 | 10.4207
| Std Dev. |  | 0.025386 | 0.20533 | 0.08165 | 0.048792 | 0.17785 | 0.085157 |  | 0.056914 | 0.051349 | 0.18777 | 0.20445 | 0.08165 | 0.12111 | 0.10389 | 0.30935 | 0.24396
| Rel.Std.Dev. |  | 0.2493% | 1.932% | 0.7752% | 0.4982% | 1.675% | 0.8344% |  | 0.5393% | 0.4984% | 1.836% | 1.988% | 0.7902% | 1.161% | 0.9746% | 2.98% | 2.363%
| PDM3 |  | -2.201% | 2.084% | 1.183% | -5.929% | 1.967% | -1.965% |  | 1.372% | -1.035% | -1.747% | -1.2% | -0.7385% | 0.2221% | 2.398% | -0.2715% | -0.8485%
| Z-Score (Absolute) |  | 0.8 | 1.55 | 1.06 | 2.85 | 1.49 | 0.67 | NA | 1.16 | 0.16 | 0.55 | 0.25 | 0 | 0.53 | 1.72 | 0.26 | 0.06
| Analytical results for As in OREAS 45f (Indicative Value 9.49 ppm) | [20 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 24 | 25 | 33 | 39 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | [10 empty] | 2
2
| 1 | 1 | 10 | < 20 | 8 |  | < 5 | < 20 | 9.3 | < 10 | < 10 | [10 empty] | 1
| 1 | 2 | 10 | < 20 | 5 |  | < 5 | < 20 | 10.3 | < 10 | < 10 | [10 empty] | 1
| 1 | 3 | 10 | < 20 | 10 |  | < 5 | < 20 | 9.8 | < 10 | < 10 | [10 empty] | 16
| 1 | 4 | 10 | < 20 | 11 |  | 5 | < 20 | 9.3 | < 10 | < 10 | [10 empty] | 9.4889
| 1 | 5 | 10 | < 20 | 9 |  | < 5 | < 20 | 8.9 | < 10 | < 10 | [10 empty] | 14
| 1 | 6 | 9 | < 20 | 8 |  | 15 | < 20 | 9 | < 10 | < 10
| Mean |  | 9.8333 |  | 8.5 |  | 10 |  | 9.4333
| Median |  | 10 |  | 8.5 |  | 10 |  | 9.3
| Std Dev. |  | 0.40825 |  | 2.0736 |  | 7.0711 |  | 0.52789
| Rel.Std.Dev. |  | 4.152% |  | 24.4% |  | 70.71% |  | 5.596%
| PDM3 |  | 3.63% |  | -10.42% |  | 5.386% |  | -0.5855%
| Z-Score (Absolute) |  | 0.57 | 0.67 | 0.3 | NA | 2.59 | 0.67 | 0.3 | 2.59 | 2.59
| Analytical results for B in OREAS 45f (Indicative Value < 20 ppm) | [20 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 12 | 24 | 33 | 39 | [15 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | [15 empty] | 1
1
| 1 | 1 | < 20 | < 50 | < 20 | < 10 | [15 empty] | 1
| 1 | 2 | 20 | < 50 | < 20 | < 10 | [15 empty] | 9
| 1 | 3 | 20 | < 50 | < 20 | 11 | [15 empty] | 16
| 1 | 4 | 20 | < 50 | < 20 | < 10 | [15 empty] | < 20
| 1 | 5 | < 20 | < 50 | < 20 | 11 | [15 empty] | 15
| 1 | 6 | < 20 | < 50 | < 20 | 10
| Mean |  | 20 |  |  | 10.6667
| Median |  | 20 |  |  | 11
| Std Dev. |  | 0 |  |  | 0.57735
| Rel.Std.Dev. |  | 0% |  |  | 5.413%
PDM3
| Z-Score (Absolute) |  | 0.47 | 2.35 | 0.47 | 0.88
| Analytical results for Ba in OREAS 45f (Certified Value 210 ppm) | [20 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 04 | 05 | 10 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  |  | ppm
|  |  | BF*OES | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*OES/MS | PF*MS | BF*OES | BF*OES | PF*MS | PF*MS | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  |  | 0
0
| 1 | 1 | 205 | 230 | 210 | 208 | 216 | 218 |  | 220 | 196 | 242 | 195 | 215.6761 | 209 | 205 | 209 | 197 | 226.6262 |  |  | 1
| 1 | 2 | 209 | 230 | 210 | 207 | 198 | 213 |  | 220 | 196 | 242 | 215 | 220.4707 | 207 | 215 | 204 | 197 | 215.1525 |  |  | #N/A
| 1 | 3 | 210 | 230 | 210 | 208 | 205 | 215 |  | 230 | 204 | 251 | 190 | 207.9944 | 206 | 210 | 176 | 215 | 213.1275 |  |  | 16
| 1 | 4 | 207 | 230 | 210 | 214 | 205 | 219 |  | 230 | 183 | 242 | 200 | 213.353 | 208 | 220 | 194 | 215 | 219.9575 |  |  | 210.4902
| 1 | 5 | 221 | 220 | 220 | 214 | 201 | 214 |  | 210 | 196 | 242 | 205 | 208.523 | 216 | 210 | 207 | 197 | 217.2113 |  |  | 35
| 1 | 6 | 210 | 220 | 210 | 210 | 204 | 215 |  | 220 | 196 | 242 | 189 | 216.8569 | 213 | 210 | 227 | 197 | 210.0725
| Mean |  | 210.3333 | 226.6667 | 211.6667 | 210.1667 | 204.8333 | 215.6667 |  | 221.6667 | 195.1667 | 243.5 | 199 | 213.8123 | 209.8333 | 211.6667 | 202.8333 | 203 | 217.0246
| Median |  | 209.5 | 230 | 210 | 209 | 204.5 | 215 |  | 220 | 196 | 242 | 197.5 | 214.5146 | 208.5 | 210 | 205.5 | 197 | 216.1819
| Std Dev. |  | 5.5737 | 5.164 | 4.0825 | 3.1252 | 6.1128 | 2.3381 |  | 7.5277 | 6.7651 | 3.6742 | 9.8995 | 4.8809 | 3.8687 | 5.164 | 16.9637 | 9.2952 | 5.794
| Rel.Std.Dev. |  | 2.65% | 2.278% | 1.929% | 1.487% | 2.984% | 1.084% |  | 3.396% | 3.466% | 1.509% | 4.975% | 2.283% | 1.844% | 2.44% | 8.363% | 4.579% | 2.67%
| PDM3 |  | -0.0745% | 7.685% | 0.5589% | -0.1537% | -2.687% | 2.459% |  | 5.31% | -7.28% | 15.68% | -5.459% | 1.578% | -0.312% | 0.5589% | -3.638% | -3.558% | 3.104%
| Z-Score (Absolute) |  | 0.07 | 1.73 | 0.07 | 0.09 | 0.68 | 0.52 | NA | 1.18 | 1.75 | 3.58 | 1.33 | 0.31 | 0.13 | 0.07 | 0.9 | 0.88 | 0.67
| Analytical results for Be in OREAS 45f (Certified Value 1.03 ppm) | [20 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 39 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*OES | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [7 empty] | 2
2
| 1 | 1 | < 5 | 1 | 1.6 |  | 1 | < 1 | 1 | 2 | 1.241 | < 5 | 0.9 | 0.87 | [7 empty] | 1
| 1 | 2 | < 5 | 1 | 1.4 |  | 2 | 1 | 1 | < 1 | 1.3212 | < 5 | 0.9 | 1.06 | [7 empty] | #N/A
| 1 | 3 | < 5 | 1 | 0.9 |  | 2 | < 1 | 1 | 1 | 1.0698 | < 5 | 0.8 | 0.93 | [7 empty] | 16
| 1 | 4 | < 5 | 1 | 1.1 |  | 2 | 2 | 1 | 1 | 1.0978 | < 5 | 0.8 | 1.05 | [7 empty] | 1.0329
| 1 | 5 | < 5 | 2 | 1.2 |  | 2 | < 1 | 1 | 1 | 1.1598 | < 5 | 0.8 | 1.01 | [7 empty] | 36
| 1 | 6 | < 5 | 1 | 1.1 |  | 2 | < 1 | 1 | 2 | 1.2579 | < 5 | 0.9 | 0.98
| Mean |  |  | 1.1667 | 1.2167 |  | 1.8333 | 1.5 | 1 | 1.4 | 1.1912 |  | 0.85 | 0.98333
| Median |  |  | 1 | 1.15 |  | 2 | 1.5 | 1 | 1 | 1.2004 |  | 0.85 | 0.995
| Std Dev. |  |  | 0.40825 | 0.24833 |  | 0.40825 | 0.70711 | 0 | 0.54772 | 0.098244 |  | 0.054772 | 0.073121
| Rel.Std.Dev. |  |  | 34.99% | 20.41% |  | 22.27% | 47.14% | 0% | 39.12% | 8.247% |  | 6.444% | 7.436%
| PDM3 |  |  | 12.95% | 17.79% |  | 77.49% | 45.22% | -3.187% | 35.54% | 15.33% |  | -17.71% | -4.8%
| Z-Score (Absolute) |  | 4.24 | NA | 0.08 | NA | NA | NA | 0.62 | NA | 0 | 4.24 | 1.11 | 0.67
Results from laboratories 13, 19, 20 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 45f (Certified Value 0.2 ppm) | [20 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 23 | 24 | 25 | 33 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.2 | 0.2 | 0.2 |  | < 1 | 0.3 | 0.1989 | 0.227 | < 0.5 | < 1 | [9 empty] | 1
| 1 | 2 | 0.2 | 0.2 | 0.2 |  | < 1 | 0.2 | 0.21921 | 0.222 | < 0.5 | < 1 | [9 empty] | #N/A
| 1 | 3 | 0.2 | 0.2 | 0.2 |  | < 1 | 0.1 | 0.22121 | 0.224 | < 0.5 | < 1 | [9 empty] | 16
| 1 | 4 | 0.2 | 0.2 | 0.3 |  | < 1 | 0.1 | 0.19097 | 0.237 | < 0.5 | < 1 | [9 empty] | 0.19993
| 1 | 5 | 0.2 | 0.2 | 0.3 |  | < 1 | 0.1 | 0.19097 | 0.216 | < 0.5 | < 1 | [9 empty] | 37
| 1 | 6 | 0.2 | 0.2 | 0.4 |  | < 1 | 0.2 | 0.2346 | 0.216 | < 0.5 | < 1
| Mean |  | 0.2 | 0.2 | 0.26667 |  |  | 0.16667 | 0.20931 | 0.22367
| Median |  | 0.2 | 0.2 | 0.25 |  |  | 0.15 | 0.20906 | 0.223
| Std Dev. |  | 3.0405e-17 | 3.0405e-17 | 0.08165 |  |  | 0.08165 | 0.018224 | 0.0078655
| Rel.Std.Dev. |  | 1.52e-14% | 1.52e-14% | 30.62% |  |  | 48.99% | 8.707% | 3.517%
| PDM3 |  | 0.03577% | 0.03577% | 33.38% |  |  | -16.64% | 4.692% | 11.87%
| Z-Score (Absolute) |  | 0.61 | 0.61 | 1.1 | NA | 7.08 | 1.46 | 0.37 | 0 | 0.67 | 7.08
| Analytical results for Ca in OREAS 45f (Certified Value 0.102 wt.%) | [20 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES | BF*OES/MS | BF*OES | BF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  | 3
3
| 1 | 1 | < 0.14 | 0.129 | 0.1 | 0.08791 | 0.1 | 0.09 | 0.1 | 0.12 |  | 0.114 | 0.1 | < 0.1 | 0.22258 | 0.102 | < 0.2 | 0.079 | 0.1001 | 0.0907 |  | 1
| 1 | 2 | < 0.14 | 0.121 | < 0.1 | 0.08969 | 0.1 | 0.1 | 0.1 | 0.11 |  | 0.1 | 0.1 | 0.1 | 0.23496 | 0.122 | < 0.2 | 0.071 | 0.1008 | 0.0964 |  | #N/A
| 1 | 3 | < 0.14 | 0.129 | < 0.1 | 0.08791 | 0.1 | 0.1 | 0.1 | 0.1 |  | 0.1 | 0.1 | < 0.1 | 0.1947 | 0.144 | < 0.2 | 0.093 | 0.1058 | 0.0943 |  | 16
| 1 | 4 | < 0.14 | 0.129 | 0.1 | 0.09077 | 0.1 | 0.09 | 0.1 | 0.1 |  | 0.1 | 0.1 | 0.1 | 0.22962 | 0.113 | < 0.2 | 0.079 | 0.1058 | 0.098 |  | 0.10244
| 1 | 5 | < 0.14 | 0.136 | < 0.1 | 0.09327 | 0.1 | 0.1 | 0.1 | 0.1 |  | 0.1 | 0.107 | 0.1 | 0.19093 | 0.126 | < 0.2 | 0.107 | 0.0993 | 0.0967 |  | 38
| 1 | 6 | < 0.14 | 0.121 | 0.1 | 0.08862 | 0.1 | 0.09 | 0.1 | 0.11 |  | 0.1 | 0.1 | 0.1 | 0.2202 | 0.115 | < 0.2 | 0.071 | 0.1015 | 0.0999
| Mean |  |  | 0.1275 | 0.1 | 0.089695 | 0.1 | 0.095 | 0.1 | 0.10667 |  | 0.10233 | 0.10117 | 0.1 | 0.2155 | 0.12033 |  | 0.083333 | 0.10222 | 0.096
| Median |  |  | 0.129 | 0.1 | 0.089155 | 0.1 | 0.095 | 0.1 | 0.105 |  | 0.1 | 0.1 | 0.1 | 0.22139 | 0.1185 |  | 0.079 | 0.10115 | 0.09655
| Std Dev. |  |  | 0.0057184 | 1.6997e-17 | 0.0020719 | 1.5202e-17 | 0.0054772 | 1.5202e-17 | 0.008165 |  | 0.0057155 | 0.0028577 | 0 | 0.018368 | 0.014236 |  | 0.014109 | 0.0028701 | 0.0031887
| Rel.Std.Dev. |  |  | 4.485% | 1.7e-14% | 2.31% | 1.52e-14% | 5.766% | 1.52e-14% | 7.655% |  | 5.585% | 2.825% | 0% | 8.523% | 11.83% |  | 16.93% | 2.808% | 3.322%
| PDM3 |  |  | 24.46% | -2.384% | -12.44% | -2.384% | -7.264% | -2.384% | 4.124% |  | -0.1059% | -1.245% | -2.384% | 110.4% | 17.47% |  | -18.65% | -0.2198% | -6.288%
| Z-Score (Absolute) |  | 3.21 | 2.83 | NA | 1.25 | NA | 0.67 | NA | 0.58 | NA | 0.13 | 0 | NA | 12.33 | 2.06 | 0.13 | 1.93 | 0.1 | 0.57
Results from laboratories 5, 12, 14 & 23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 45f (Indicative Value < 0.2 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 24 | 25 | 33 | 39 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | [10 empty] | 2
2
| 1 | 1 | < 0.2 | < 10 | < 0.8 |  | < 10 | < 10 | < 0.2 | < 0.5 | < 0.5 | [10 empty] | 1
| 1 | 2 | < 0.2 | < 10 | < 0.8 |  | < 10 | < 10 | < 0.2 | < 0.5 | < 0.5 | [10 empty] | 10
| 1 | 3 | < 0.2 | < 10 | < 0.8 |  | < 10 | < 10 | < 0.2 | < 0.5 | < 0.5 | [10 empty] | 16
| 1 | 4 | < 0.2 | < 10 | < 0.8 |  | < 10 | < 10 | < 0.2 | < 0.5 | < 0.5 | [10 empty] | < 0.2
| 1 | 5 | < 0.2 | < 10 | < 0.8 |  | < 10 | < 10 | < 0.2 | < 0.5 | < 0.5 | [10 empty] | 16
| 1 | 6 | < 0.2 | < 10 | 0.9 |  | < 10 | < 10 | < 0.2 | < 0.5 | < 0.5
| Mean |  |  |  | 0.9
| Median |  |  |  | 0.9
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 0.67 | 11.72 | 0.3 | NA | 11.72 | 11.72 | 0.67 | 0.3 | 0.3
| Analytical results for Ce in OREAS 45f (Certified Value 29.4 ppm) | [20 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 1
2
| 1 | 1 | 30.4 | 47.9237 | 28.9 | 30 | 29 | 28.2 | 30.9 | 29.2 | 30.5 | 46 | [9 empty] | 1
| 1 | 2 | 31.4 | 48.4128 | 27.2 | 28.5 | 28.5 | 28.2 | 30.3 | 28.7 | 31.5 | 38 | [9 empty] | #N/A
| 1 | 3 | 30.5 | 47.0781 | 26.9 | 32 | 28.6 | 27.8 | 32.9 | 28 | 32.1 | 30.83 | [9 empty] | 16
| 1 | 4 | 30.8 | 46.8463 | 28 | 31 | 28.8 | 28.2 | 31.4 | 27.1 | 30.5 | 44.945 | [9 empty] | 29.385
| 1 | 5 | 30.2 | 46.3015 | 27.5 | 31 | 25.8 | 27.7 | 31 | 28.7 | 30.1 | 39.237 | [9 empty] | 39
| 1 | 6 | 30.1 | 46.8097 | 27.7 | 30 | 28.9 | 23.6 | 30.3 | 25.9 | 28.6 | 44.713
| Mean |  | 30.5667 | 47.2287 | 27.7 | 30.4167 | 28.2667 | 27.2833 | 31.1333 | 27.9333 | 30.55 | 40.6208
| Median |  | 30.45 | 46.9622 | 27.6 | 30.5 | 28.7 | 28 | 30.95 | 28.35 | 30.5 | 41.975
| Std Dev. |  | 0.4761 | 0.78603 | 0.70143 | 1.2007 | 1.2226 | 1.8181 | 0.96471 | 1.234 | 1.2095 | 5.8142
| Rel.Std.Dev. |  | 1.558% | 1.664% | 2.532% | 3.947% | 4.325% | 6.664% | 3.099% | 4.418% | 3.959% | 14.31%
| PDM3 |  | 4.021% | 60.72% | -5.734% | 3.511% | -3.806% | -7.152% | 5.95% | -4.94% | 3.965% | 38.24%
| Z-Score (Absolute) |  | 0.02 | 4.74 | 0.79 | 0.02 | 0.63 | 0.91 | 0.18 | 0.72 | 0.02 | 2.87
| Analytical results for Co in OREAS 45f (Certified Value 49 ppm) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 05 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES |  |  |  |  | 1
2
| 1 | 1 | 45 | 50 | 60 | 46 | 49.6 |  | 40 | 41.8 | 51 | 38 | 49.7264 | 49.6 | 51 | 47.4 | 55.3254 |  |  |  |  | 1
| 1 | 2 | 45.4 | 54.1667 | 40 | 46 | 50.3 |  | 40 | 39.9 | 50 | 39 | 50.9297 | 50.8 | 49 | 44.9 | 53.144 |  |  |  |  | 3
| 1 | 3 | 46 | 54.1667 | 60 | 43 | 48.5 |  | 50 | 42.2 | 49 | 37 | 45.9203 | 49.8 | 47 | 46.8 | 57.0999 |  |  |  |  | 16
| 1 | 4 | 46.2 | 50 | 60 | 48 | 49 |  | 50 | 41.5 | 51 | 39 | 46.8106 | 50.7 | 47 | 44.9 | 54.9666 |  |  |  |  | 49.0231
| 1 | 5 | 44.3 | 45.8333 | 40 | 47 | 47.7 |  | 50 | 39.9 | 50 | 40 | 47.5284 | 49.3 | 49 | 47.7 | 55.0121 |  |  |  |  | 40
| 1 | 6 | 45.1 | 54.1667 | 60 | 47 | 48.9 |  | 40 | 39 | 50 | 40 | 48.1347 | 49.4 | 49 | 49.4 | 51.7569
| Mean |  | 45.3333 | 51.3889 | 53.3333 | 46.1667 | 49 |  | 45 | 40.7167 | 50.1667 | 38.8333 | 48.175 | 49.9333 | 48.6667 | 46.85 | 54.5508
| Median |  | 45.25 | 52.0833 | 60 | 46.5 | 48.95 |  | 45 | 40.7 | 50 | 39 | 47.8315 | 49.7 | 49 | 47.1 | 54.9894
| Std Dev. |  | 0.69761 | 3.4021 | 10.328 | 1.7224 | 0.89443 |  | 5.4772 | 1.2859 | 0.75277 | 1.169 | 1.8633 | 0.65625 | 1.5055 | 1.7398 | 1.8584
| Rel.Std.Dev. |  | 1.539% | 6.62% | 19.36% | 3.731% | 1.825% |  | 12.17% | 3.158% | 1.501% | 3.01% | 3.868% | 1.314% | 3.094% | 3.714% | 3.407%
| PDM3 |  | -7.527% | 4.826% | 8.792% | -5.827% | -0.0472% |  | -8.207% | -16.94% | 2.333% | -20.79% | -1.73% | 1.857% | -0.7271% | -4.433% | 11.28%
| Z-Score (Absolute) |  | 1.04 | 1 | NA | 0.76 | 0.2 | NA | NA | 2.6 | 0.59 | 3.23 | 0.08 | 0.51 | 0.08 | 0.53 | 2.07
Results from laboratories 12 & 19 were removed due to their 10 ppm reading resolution.
| Analytical results for Cr in OREAS 45f (Certified Value 476 ppm) | [20 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 04 | 05 | 10 | 12 | 13 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  |  |  | ppm
|  |  | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*OES | BF*OES | BF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  |  |  | 0
0
| 1 | 1 | 410 | 480 | 480 | 450 | 500 |  | 470 | 479 | 452 | 503 | 483.6579 | 461 | 535 | 410 | 445 | 427 |  |  |  | 1
| 1 | 2 | 410 | 500 | 470 | 450 | 500 |  | 450 | 472 | 445 | 492 | 491.1526 | 502 | 510 | 410 | 458 | 398 |  |  |  | #N/A
| 1 | 3 | 480 | 490 | 485 | 450 | 500 |  | 460 | 472 | 424 | 465 | 461.8704 | 472 | 515 | 480 | 479 | 437 |  |  |  | 16
| 1 | 4 | 410 | 470 | 485 | 450 | 500 |  | 460 | 472 | 445 | 522 | 465.5302 | 466 | 515 | 410 | 479 | 405 |  |  |  | 475.5643
| 1 | 5 | 410 | 490 | 565 | 450 | 500 |  | 470 | 479 | 452 | 483 | 468.4129 | 484 | 500 | 480 | 452 | 421 |  |  |  | 41
| 1 | 6 | 480 | 480 | 540 | 450 | 500 |  | 460 | 472 | 445 | 513 | 474.3497 | 473 | 490 | 410 | 465 | 399
| Mean |  | 433.3333 | 485 | 504.1667 | 450 | 500 |  | 461.6667 | 474.3333 | 443.8333 | 496.3333 | 474.1623 | 476.3333 | 510.8333 | 433.3333 | 463 | 414.5
| Median |  | 410 | 485 | 485 | 450 | 500 |  | 460 | 472 | 445 | 497.5 | 471.3813 | 472.5 | 512.5 | 410 | 461.5 | 413
| Std Dev. |  | 36.1478 | 10.4881 | 38.6545 | 0 | 0 |  | 7.5277 | 3.6148 | 10.3037 | 20.7814 | 11.2919 | 14.7603 | 15.3025 | 36.1478 | 14.0428 | 16.171
| Rel.Std.Dev. |  | 8.342% | 2.162% | 7.667% | 0% | 0% |  | 1.631% | 0.7621% | 2.322% | 4.187% | 2.381% | 3.099% | 2.996% | 8.342% | 3.033% | 3.901%
| PDM3 |  | -8.88% | 1.984% | 6.014% | -5.376% | 5.138% |  | -2.922% | -0.2588% | -6.672% | 4.367% | -0.2948% | 0.1617% | 7.416% | -8.88% | -2.642% | -12.84%
| Z-Score (Absolute) |  | NA | 0.42 | 1.16 | 0.94 | NA | NA | 0.49 | 0 | 1.19 | 0.86 | 0 | 0.08 | 1.42 | NA | 0.44 | 2.33
Results from laboratories 4, 13 & 34 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 45f (Certified Value 3.66 ppm) | [20 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 10 | 13 | 14 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 3.7 | 5.7253 | 3.53 | 3.9 | 3 | 3.4 | 3.5 | 3.4 | 4.0924 | 3.99 | 3.6 | 3.77 | 3.656 | [6 empty] | 1
| 1 | 2 | 3.7 | 5.8417 | 3.59 | 3.4 | 3 | 3 | 3.2 | 3.1 | 4.0431 | 3.87 | 3.6 | 3.67 | 3.689 | [6 empty] | #N/A
| 1 | 3 | 3.8 | 5.8731 | 3.53 | 3.7 | 3 | 3.3 | 3.6 | 3.9 | 3.6147 | 4.12 | 3.6 | 3.89 | 3.341 | [6 empty] | 16
| 1 | 4 | 3.7 | 5.6423 | 3.47 | 3.8 | 3 | 3.3 | 3.5 | 3.8 | 3.7306 | 3.92 | 3.8 | 3.77 | 3.411 | [6 empty] | 3.6605
| 1 | 5 | 3.7 | 4.6906 | 3.56 | 3.6 | 3 | 3.4 | 3.7 | 3.3 | 3.799 | 4.18 | 3.6 | 3.95 | 3.742 | [6 empty] | 42
| 1 | 6 | 3.7 | 5.6151 | 3.75 | 4.1 | 3 | 3.4 | 3.3 | 3.5 | 4.1658 | 3.9 | 4.2 | 3.82 | 3.429
| Mean |  | 3.7167 | 5.5647 | 3.5717 | 3.75 | 3 | 3.3 | 3.4667 | 3.5 | 3.9076 | 3.9967 | 3.7333 | 3.8117 | 3.5447
| Median |  | 3.7 | 5.6838 | 3.545 | 3.75 | 3 | 3.35 | 3.5 | 3.45 | 3.921 | 3.955 | 3.6 | 3.795 | 3.5425
| Std Dev. |  | 0.040825 | 0.44049 | 0.096003 | 0.2429 | 0 | 0.15492 | 0.18619 | 0.30332 | 0.22277 | 0.1266 | 0.24221 | 0.098877 | 0.17023
| Rel.Std.Dev. |  | 1.098% | 7.916% | 2.688% | 6.477% | 0% | 4.695% | 5.371% | 8.666% | 5.701% | 3.168% | 6.488% | 2.594% | 4.802%
| PDM3 |  | 1.534% | 52.02% | -2.427% | 2.445% | -18.04% | -9.849% | -5.295% | -4.385% | 6.75% | 9.183% | 1.989% | 4.129% | -3.165%
| Z-Score (Absolute) |  | 0.03 | 6.84 | 0.57 | 0.09 | NA | 1.58 | 0.96 | 0.84 | 0.68 | 1.01 | 0.03 | 0.32 | 0.67
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 45f (Certified Value 362 ppm) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 10 | 12 | 13 | 14 | 17 | 19 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*MS | PF*OES | [6 empty] | 0
0
| 1 | 1 | 350 | 660 | 400 | 300 | 390 |  | 360 | 354 | 354.0925 | 352 | 370 | 387 | 363 | [6 empty] | 1
| 1 | 2 | 350 | 355 | 350 | 300 | 380 |  | 360 | 363 | 353.2111 | 362 | 360 | 378 | 342 | [6 empty] | #N/A
| 1 | 3 | 360 | 360 | 400 | 300 | 380 |  | 380 | 328 | 336.6577 | 356 | 355 | 388 | 357 | [6 empty] | 16
| 1 | 4 | 350 | 370 | 350 | 300 | 390 |  | 360 | 353 | 364.8101 | 358 | 355 | 375 | 350 | [6 empty] | 361.5044
| 1 | 5 | 350 | 370 | 350 | 300 | 370 |  | 360 | 349 | 355.3889 | 356 | 365 | 377 | 359 | [6 empty] | 43
| 1 | 6 | 340 | 360 | 350 | 300 | 390 |  | 360 | 345 | 343.1288 | 353 | 355 | 372 | 357
| Mean |  | 350 | 412.5 | 366.6667 | 300 | 383.3333 |  | 363.3333 | 348.6667 | 351.2148 | 356.1667 | 360 | 379.5 | 354.6667
| Median |  | 350 | 365 | 350 | 300 | 385 |  | 360 | 351 | 353.6518 | 356 | 357.5 | 377.5 | 357
| Std Dev. |  | 6.3246 | 121.3981 | 25.8199 | 0 | 8.165 |  | 8.165 | 11.7757 | 9.9178 | 3.6009 | 6.3246 | 6.5345 | 7.5011
| Rel.Std.Dev. |  | 1.807% | 29.43% | 7.042% | 0% | 2.13% |  | 2.247% | 3.377% | 2.824% | 1.011% | 1.757% | 1.722% | 2.115%
| PDM3 |  | -3.182% | 14.11% | 1.428% | -17.01% | 6.038% |  | 0.5059% | -3.551% | -2.846% | -1.477% | -0.4161% | 4.978% | -1.891%
| Z-Score (Absolute) |  | 0.65 | 4.4 | 0.69 | 4.7 | 2.04 | NA | 0.42 | 0.76 | 0.56 | 0.16 | 0.16 | 1.73 | 0.28
| Analytical results for Dy in OREAS 45f (Certified Value 2.88 ppm) | [20 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 2
3
| 1 | 1 | 2.6 | 3.0876 | 3.16 | 2.5 | 2.5 | 2.99 | 3.03 | 2.2 | 2.92 | 1.391 | [9 empty] | 1
| 1 | 2 | 2.59 | 3.1301 | 3.19 | 2.5 | 2.9 | 3.08 | 2.87 | 2.2 | 3.02 | 1.174 | [9 empty] | #N/A
| 1 | 3 | 2.65 | 3.0555 | 2.69 | 3 | 2.69 | 3.13 | 2.81 | 2.1 | 2.95 | 2.37 | [9 empty] | 16
| 1 | 4 | 2.71 | 3.0594 | 3.12 | 2.5 | 2.69 | 3.01 | 3.07 | 2.1 | 2.93 | 1.324 | [9 empty] | 2.8793
| 1 | 5 | 2.76 | 2.9635 | 2.74 | 3 | 2.46 | 2.97 | 2.9 | 2.1 | 2.78 | 1.012 | [9 empty] | 44
| 1 | 6 | 2.69 | 2.9737 | 2.81 | 2.5 | 2.59 | 2.99 | 2.95 | 2.2 | 2.72 | 2.163
| Mean |  | 2.6667 | 3.045 | 2.9517 | 2.6667 | 2.6383 | 3.0283 | 2.9383 | 2.15 | 2.8867 | 1.5723
| Median |  | 2.67 | 3.0575 | 2.965 | 2.5 | 2.64 | 3 | 2.925 | 2.15 | 2.925 | 1.3575
| Std Dev. |  | 0.065929 | 0.064967 | 0.22886 | 0.2582 | 0.15943 | 0.062743 | 0.098472 | 0.054772 | 0.11308 | 0.5572
| Rel.Std.Dev. |  | 2.472% | 2.134% | 7.754% | 9.682% | 6.043% | 2.072% | 3.351% | 2.548% | 3.917% | 35.44%
| PDM3 |  | -7.384% | 5.755% | 2.514% | -7.384% | -8.368% | 5.177% | 2.051% | -25.33% | 0.2565% | -45.39%
| Z-Score (Absolute) |  | 0.94 | 0.67 | 0.28 | NA | 1.06 | 0.6 | 0.22 | 3.14 | 0 | 5.6
| Analytical results for Er in OREAS 45f (Certified Value 1.65 ppm) | [20 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 1.69 | 1.8774 | 2.02 | 1.5 | 1.84 | 1.45 | 1.8 | 1.5 | 1.79 | 1.62 | [9 empty] | 1
| 1 | 2 | 1.71 | 1.9 | 1.84 | 1.5 | 1.8 | 1.39 | 1.88 | 1.4 | 1.73 | 1.602 | [9 empty] | #N/A
| 1 | 3 | 1.6 | 1.8056 | 1.56 | 1.5 | 1.84 | 1.44 | 1.82 | 1.4 | 1.79 | 1.445 | [9 empty] | 16
| 1 | 4 | 1.69 | 1.8299 | 1.68 | 1.5 | 1.77 | 1.46 | 1.82 | 1.3 | 1.81 | 1.467 | [9 empty] | 1.6499
| 1 | 5 | 1.68 | 1.7717 | 1.63 | 2 | 1.79 | 1.47 | 1.77 | 1.3 | 1.84 | 1.586 | [9 empty] | 45
| 1 | 6 | 1.56 | 1.7502 | 1.81 | 1.5 | 1.76 | 1.49 | 1.78 | 1.4 | 1.75 | 1.49
| Mean |  | 1.655 | 1.8225 | 1.7567 | 1.5833 | 1.8 | 1.45 | 1.8117 | 1.3833 | 1.785 | 1.535
| Median |  | 1.685 | 1.8178 | 1.745 | 1.5 | 1.795 | 1.455 | 1.81 | 1.4 | 1.79 | 1.538
| Std Dev. |  | 0.060249 | 0.058586 | 0.16717 | 0.20412 | 0.034059 | 0.034059 | 0.0392 | 0.075277 | 0.039875 | 0.076242
| Rel.Std.Dev. |  | 3.64% | 3.215% | 9.516% | 12.89% | 1.892% | 2.349% | 2.164% | 5.442% | 2.234% | 4.967%
| PDM3 |  | 0.3083% | 10.46% | 6.47% | -4.035% | 9.097% | -12.12% | 9.804% | -16.16% | 8.187% | -6.965%
| Z-Score (Absolute) |  | 0.31 | 0.71 | 0.31 | 0.74 | 0.57 | 1.55 | 0.64 | 1.96 | 0.48 | 1.04
| Analytical results for Eu in OREAS 45f (Certified Value 0.62 ppm) | [20 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.67 | 1.09 | 0.7 | 0.6 | 0.6 | 0.52 | 0.711 | 0.6 | 0.65 | 0.625 | [9 empty] | 1
| 1 | 2 | 0.7 | 1.1193 | 0.71 | 0.6 | 0.64 | 0.53 | 0.701 | 0.6 | 0.71 | 0.631 | [9 empty] | #N/A
| 1 | 3 | 0.63 | 1.0427 | 0.45 | 0.6 | 0.64 | 0.54 | 0.748 | 0.5 | 0.67 | 0.567 | [9 empty] | 16
| 1 | 4 | 0.58 | 1.0745 | 0.51 | 0.6 | 0.65 | 0.55 | 0.708 | 0.5 | 0.69 | 0.583 | [9 empty] | 0.62283
| 1 | 5 | 0.62 | 1.0748 | 0.58 | 0.6 | 0.6 | 0.55 | 0.719 | 0.6 | 0.61 | 0.681 | [9 empty] | 46
| 1 | 6 | 0.57 | 1.0291 | 0.56 | 0.6 | 0.64 | 0.55 | 0.706 | 0.6 | 0.65 | 0.548
| Mean |  | 0.62833 | 1.0717 | 0.585 | 0.6 | 0.62833 | 0.54 | 0.7155 | 0.56667 | 0.66333 | 0.60583
| Median |  | 0.625 | 1.0746 | 0.57 | 0.6 | 0.64 | 0.545 | 0.7095 | 0.6 | 0.66 | 0.604
| Std Dev. |  | 0.050365 | 0.032453 | 0.1033 | 0 | 0.022286 | 0.012649 | 0.017003 | 0.05164 | 0.035024 | 0.04904
| Rel.Std.Dev. |  | 8.016% | 3.028% | 17.66% | 0% | 3.547% | 2.342% | 2.376% | 9.113% | 5.28% | 8.095%
| PDM3 |  | 0.8831% | 72.07% | -6.074% | -3.666% | 0.8831% | -13.3% | 14.88% | -9.018% | 6.503% | -2.729%
| Z-Score (Absolute) |  | 0 | 7.63 | 0.75 | NA | 0 | 1.52 | 1.5 | NA | 0.6 | 0.39
Results from laboratory 19 & 33 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 45f (Certified Value 15.16 wt.%) | [20 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES | BF*OES/MS | PF*OES | BF*OES | BF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES | 2
3
| 1 | 1 | 14.0392 | 14.828 | 16 | 14.7136 | 15.5 | 15.03 | 15.25 | 15.61 |  | 15.1 | 14.947 | 15.849 | 14.83 | 15.2339 | 14.8 | 15 | 15.087 | 15.178 | 15.3152 | 1
| 1 | 2 | 14.0017 | 15.143 | 15.5 | 14.774 | 15.5 | 15.41 | 14.95 | 15.51 |  | 15.2 | 15.073 | 15.982 | 15.18 | 15.0683 | 14.9 | 15.4 | 14.542 | 15.178 | 15.0329 | #N/A
| 1 | 3 | 14.0644 | 14.863 | 16 | 14.7673 | 15.2 | 14.95 | 15 | 15.64 |  | 15 | 14.94 | 15.766 | 14.76 | 15.0409 | 14.9 | 15.1 | 15.003 | 15.9474 | 15.2315 | 16
| 1 | 4 | 13.9777 | 14.94 | 15.7 | 14.7908 | 15.5 | 14.97 | 15.25 | 15.84 |  | 15 | 15.094 | 16.024 | 15.02 | 14.9908 | 14.9 | 15.6 | 14.716 | 16.1572 | 15.4092 | 15.1571
| 1 | 5 | 14.0804 | 14.954 | 15.5 | 14.8074 | 15.4 | 15.58 | 14.65 | 15.58 |  | 15.2 | 15.08 | 16.227 | 15.23 | 15.1908 | 14.9 | 15 | 14.653 | 15.3878 | 15.4963 | 47
| 1 | 6 | 14.0782 | 14.961 | 15.5 | 14.8412 | 15.6 | 14.93 | 15.35 | 15.92 |  | 14.9 | 14.898 | 16.073 | 14.37 | 15.0433 | 14.8 | 15.5 | 14.611 | 15.3878 | 14.8479
| Mean |  | 14.0403 | 14.9482 | 15.7 | 14.7824 | 15.45 | 15.145 | 15.075 | 15.6833 |  | 15.0667 | 15.0053 | 15.9868 | 14.8983 | 15.0947 | 14.8667 | 15.2667 | 14.7687 | 15.5394 | 15.2222
| Median |  | 14.0518 | 14.947 | 15.6 | 14.7824 | 15.5 | 15 | 15.125 | 15.625 |  | 15.05 | 15.01 | 16.003 | 14.925 | 15.0558 | 14.9 | 15.25 | 14.6845 | 15.3878 | 15.2734
| Std Dev. |  | 0.042504 | 0.10947 | 0.24495 | 0.042878 | 0.13784 | 0.27841 | 0.26029 | 0.16033 |  | 0.12111 | 0.086264 | 0.16387 | 0.31846 | 0.095572 | 0.05164 | 0.26583 | 0.223 | 0.4136 | 0.24281
| Rel.Std.Dev. |  | 0.3027% | 0.7323% | 1.56% | 0.2901% | 0.8922% | 1.838% | 1.727% | 1.022% |  | 0.8038% | 0.5749% | 1.025% | 2.138% | 0.6332% | 0.3474% | 1.741% | 1.51% | 2.662% | 1.595%
| PDM3 |  | -7.368% | -1.378% | 3.582% | -2.472% | 1.933% | -0.07956% | -0.5414% | 3.472% |  | -0.5964% | -1.001% | 5.475% | -1.707% | -0.4117% | -1.916% | 0.7231% | -2.562% | 2.522% | 0.4296%
| Z-Score (Absolute) |  | 3.48 | 0.45 | 2.05 | 1.01 | 1.22 | 0.2 | 0.03 | 1.99 | NA | 0.06 | 0.26 | 3.01 | 0.62 | 0.03 | 0.73 | 0.61 | 1.05 | 1.51 | 0.46
| Analytical results for Ga in OREAS 45f (Certified Value 26.2 ppm) | [20 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [7 empty] | 1
2
| 1 | 1 | 27 | 27 | 27.6 |  | 23.2 | 23.6 | 23 | 27.0505 | 30.3 | 27 | 27.3 | 23.506 | [7 empty] | 1
| 1 | 2 | 27 | 27 | 29.4 |  | 24.1 | 23.9 | 25 | 24.8346 | 29.8 | 27 | 27.3 | 24.23 | [7 empty] | #N/A
| 1 | 3 | 28 | 26 | 27.7 |  | 22.8 | 24.2 | 24 | 25.459 | 29.8 | 25 | 26.7 | 21.822 | [7 empty] | 16
| 1 | 4 | 27 | 28 | 29.3 |  | 24.6 | 24.1 | 24 | 28.0174 | 29.7 | 25 | 27.2 | 22.464 | [7 empty] | 26.1739
| 1 | 5 | 27 | 27 | 29.4 |  | 22.3 | 24.9 | 23 | 27.9365 | 30.4 | 27 | 28.6 | 24.93 | [7 empty] | 48
| 1 | 6 | 26 | 27 | 30.4 |  | 20.7 | 25 | 24 | 27.586 | 30.6 | 27 | 28.6 | 21.144
| Mean |  | 27 | 27 | 28.9667 |  | 22.95 | 24.2833 | 23.8333 | 26.814 | 30.1 | 26.3333 | 27.6167 | 23.016
| Median |  | 27 | 27 | 29.35 |  | 23 | 24.15 | 24 | 27.3182 | 30.05 | 27 | 27.3 | 22.985
| Std Dev. |  | 0.63246 | 0.63246 | 1.0967 |  | 1.3867 | 0.55648 | 0.75277 | 1.3501 | 0.37947 | 1.0328 | 0.79352 | 1.4568
| Rel.Std.Dev. |  | 2.342% | 2.342% | 3.786% |  | 6.042% | 2.292% | 3.158% | 5.035% | 1.261% | 3.922% | 2.873% | 6.33%
| PDM3 |  | 3.156% | 3.156% | 10.67% |  | -12.32% | -7.223% | -8.943% | 2.445% | 15% | 0.609% | 5.512% | -12.07%
| Z-Score (Absolute) |  | 0.06 | 0.06 | 0.67 | NA | 1.21 | 0.79 | 0.93 | 0 | 1.03 | 0.15 | 0.25 | 1.19
| Analytical results for Gd in OREAS 45f (Certified Value 2.57 ppm) | [20 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 2
2
| 1 | 1 | 2.5 | 2.5324 | 2.58 | 2 | 2.31 | 2.96 | 2.82 | 2.3 | 2.65 | < 2 | [9 empty] | 1
| 1 | 2 | 2.38 | 2.5627 | 2.71 | 2 | 2.27 | 2.78 | 2.84 | 2.3 | 2.77 | < 2 | [9 empty] | #N/A
| 1 | 3 | 2.54 | 2.436 | 2.63 | 2 | 2.37 | 2.97 | 2.63 | 2.1 | 2.72 | < 2 | [9 empty] | 16
| 1 | 4 | 2.59 | 2.4975 | 3 | 2 | 2.41 | 2.8 | 2.79 | 2 | 2.83 | < 2 | [9 empty] | 2.5736
| 1 | 5 | 2.5 | 2.3038 | 2.74 | 2 | 2.31 | 3.05 | 2.63 | 2.2 | 2.52 | < 2 | [9 empty] | 49
| 1 | 6 | 2.44 | 2.3611 | 2.79 | 2 | 2.41 | 3.23 | 2.71 | 2.2 | 2.56 | < 2
| Mean |  | 2.4917 | 2.4489 | 2.7417 | 2 | 2.3467 | 2.965 | 2.7367 | 2.1833 | 2.675
| Median |  | 2.5 | 2.4667 | 2.725 | 2 | 2.34 | 2.965 | 2.75 | 2.2 | 2.685
| Std Dev. |  | 0.073869 | 0.1012 | 0.14743 | 0 | 0.058538 | 0.1667 | 0.093737 | 0.1169 | 0.12079
| Rel.Std.Dev. |  | 2.965% | 4.132% | 5.378% | 0% | 2.495% | 5.622% | 3.425% | 5.354% | 4.515%
| PDM3 |  | -3.184% | -4.846% | 6.53% | -22.29% | -8.818% | 15.21% | 6.336% | -15.16% | 3.939%
| Z-Score (Absolute) |  | 0 | 0.12 | 0.69 | NA | 0.4 | 1.3 | 0.67 | 0.85 | 0.5 | 4.11
Results from laboratory 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 45f (Indicative Value 2.25 ppm) | [20 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 05 | 14 | 23 | 24 | 25 | 39 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | 2 | 2.4 | 2 | 1.7644 | 2.5 | 0.147 | [13 empty] | 1
| 1 | 2 | 2 | 2.4 | 2 | 1.8406 | 2.48 | 0.214 | [13 empty] | 11
| 1 | 3 | 2 | 2.3 | 2 | 1.1267 | 2.64 | 0.103 | [13 empty] | 16
| 1 | 4 | 2 | 2.6 | 2 | 2.1064 | 2.32 | < 0.1 | [13 empty] | 2.2454
| 1 | 5 | 2 | 2.5 | 2 | 1.8513 | 2.61 | 0.302 | [13 empty] | 17
| 1 | 6 | 2 | 2.8 | 2 | < 1 | 2.44 | < 0.1
| Mean |  | 2 | 2.5 | 2 | 1.7379 | 2.4983 | 0.1915
| Median |  | 2 | 2.45 | 2 | 1.8406 | 2.49 | 0.1805
| Std Dev. |  | 0 | 0.17889 | 0 | 0.36521 | 0.11669 | 0.086658
| Rel.Std.Dev. |  | 0% | 7.155% | 0% | 21.01% | 4.671% | 45.25%
| PDM3 |  | -10.93% | 11.34% | -10.93% | -22.6% | 11.26% | -91.47%
| Z-Score (Absolute) |  | NA | 0.68 | NA | 0.67 | 0.67 | 2.6
Results from laboratories 5 & 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Hf in OREAS 45f (Certified Value 6.97 ppm) | [20 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 05 | 19 | 20 | 21 | 25 | 34 | 39 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | BF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | 5 | 6 | 6.2 | 7.8 | 7.45 | 6.85 | 5.6 | [12 empty] | 1
| 1 | 2 | 6 | 4 | 6.4 | 8 | 7.27 | 6.64 | 5.8 | [12 empty] | #N/A
| 1 | 3 | 6 | 6 | 6.6 | 8.4 | 7.53 | 6.8 | 5.9 | [12 empty] | 16
| 1 | 4 | 6 | 4 | 6.6 | 7.8 | 7.62 | 7.11 | 6.2 | [12 empty] | 6.9683
| 1 | 5 | 6 | 4 | 6.6 | 8.6 | 7.35 | 7.17 | 5.9 | [12 empty] | 50
| 1 | 6 | 6 | 6 | 6.3 | 7.5 | 7.72 | 7.14 | 6.2
| Mean |  | 5.8333 | 5 | 6.45 | 8.0167 | 7.49 | 6.9517 | 5.9333
| Median |  | 6 | 5 | 6.5 | 7.9 | 7.49 | 6.98 | 5.9
| Std Dev. |  | 0.40825 | 1.0954 | 0.17607 | 0.41191 | 0.16793 | 0.21849 | 0.23381
| Rel.Std.Dev. |  | 6.999% | 21.91% | 2.73% | 5.138% | 2.242% | 3.143% | 3.941%
| PDM3 |  | -16.29% | -28.25% | -7.438% | 15.04% | 7.486% | -0.2392% | -14.85%
| Z-Score (Absolute) |  | NA | NA | 0.63 | 1.33 | 0.67 | 0 | 1.28
Results from laboratory 5 & 19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 45f (Certified Value 0.56 ppm) | [20 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.56 | 0.56241 | 0.54 | 0.6 | 0.57 | 0.59 | 0.613 | 0.5 | 0.58 | 0.34 | [9 empty] | 1
| 1 | 2 | 0.55 | 0.58138 | 0.53 | 0.6 | 0.55 | 0.66 | 0.606 | 0.5 | 0.6 | 0.33 | [9 empty] | #N/A
| 1 | 3 | 0.53 | 0.55652 | 0.47 | 0.6 | 0.58 | 0.66 | 0.588 | 0.4 | 0.61 | 0.32 | [9 empty] | 16
| 1 | 4 | 0.6 | 0.56108 | 0.48 | 0.6 | 0.64 | 0.6 | 0.594 | 0.4 | 0.64 | 0.32 | [9 empty] | 0.56015
| 1 | 5 | 0.57 | 0.53233 | 0.55 | 0.6 | 0.6 | 0.61 | 0.598 | 0.5 | 0.58 | 0.34 | [9 empty] | 51
| 1 | 6 | 0.52 | 0.53541 | 0.53 | 0.6 | 0.59 | 0.58 | 0.589 | 0.4 | 0.6 | 0.38
| Mean |  | 0.555 | 0.55485 | 0.51667 | 0.6 | 0.58833 | 0.61667 | 0.598 | 0.45 | 0.60167 | 0.33833
| Median |  | 0.555 | 0.5588 | 0.53 | 0.6 | 0.585 | 0.605 | 0.596 | 0.45 | 0.6 | 0.335
| Std Dev. |  | 0.02881 | 0.018373 | 0.033267 | 0 | 0.030605 | 0.035024 | 0.009859 | 0.054772 | 0.022286 | 0.022286
| Rel.Std.Dev. |  | 5.191% | 3.311% | 6.439% | 0% | 5.202% | 5.68% | 1.649% | 12.17% | 3.704% | 6.587%
| PDM3 |  | -0.9191% | -0.9451% | -7.763% | 7.114% | 5.032% | 10.09% | 6.757% | -19.66% | 7.412% | -39.6%
| Z-Score (Absolute) |  | 0 | 0 | 0.6 | NA | 0.52 | 0.97 | 0.67 | 1.65 | 0.73 | 3.4
Results from laboratory 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for I in OREAS 45f (Indicative Value 15.8 ppm) | [20 empty] | Indicative
I | Round | Replicate | Lab | [18 empty] | 1
| No | No | 39 | [18 empty] | ppm
|  |  | PF*MS | [18 empty] | 1
1
| 1 | 1 | 16 | [18 empty] | 1
| 1 | 2 | 15 | [18 empty] | 12
| 1 | 3 | 12 | [18 empty] | 16
| 1 | 4 | 23 | [18 empty] | 15.8333
| 1 | 5 | 13 | [18 empty] | 18
| 1 | 6 | 16
| Mean |  | 15.8333
| Median |  | 15.5
| Std Dev. |  | 3.8687
| Rel.Std.Dev. |  | 24.43%
| PDM3 |  | 2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 45f (Indicative Value < 0.2 ppm) | [20 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 05 | 13 | 14 | 19 | 24 | 25 | 33 | 39 | [11 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [11 empty] | 2
2
| 1 | 1 | < 0.2 | < 0.1 | < 0.3 | < 0.2 | < 0.1 | < 0.2 | < 0.25 | 0.125 | [11 empty] | 1
| 1 | 2 | < 0.2 | < 0.1 | < 0.3 | < 0.2 | 0.10662 | < 0.2 | < 0.25 | 0.122 | [11 empty] | 13
| 1 | 3 | < 0.2 | 0.1 | < 0.3 | < 0.2 | < 0.1 | < 0.2 | < 0.25 | 0.131 | [11 empty] | 16
| 1 | 4 | < 0.2 | < 0.1 | < 0.3 | < 0.2 | 0.12473 | < 0.2 | < 0.25 | 0.118 | [11 empty] | < 0.2
| 1 | 5 | < 0.2 | < 0.1 | < 0.3 | < 0.2 | 0.12594 | < 0.2 | < 0.25 | 0.135 | [11 empty] | 19
| 1 | 6 | < 0.2 | < 0.1 | < 0.3 | < 0.2 | 0.13504 | < 0.2 | < 0.25 | 0.109
| Mean |  |  | 0.1 |  |  | 0.12308 |  |  | 0.12333
| Median |  |  | 0.1 |  |  | 0.12533 |  |  | 0.1235
| Std Dev. | [5 empty] | 0.011902 |  |  | 0.0093095
| Rel.Std.Dev. | [5 empty] | 9.67% |  |  | 7.548%
PDM3
| Z-Score (Absolute) |  | 0 | 2.28 | 2.74 | 0 | 0.07 | 0 | 1.37 | 1.28
| Analytical results for Ir in OREAS 45f (Indicative Value < 15 ppb) | [20 empty] | Indicative
Ir | Round | Replicate | Lab | [18 empty] | 1
| No | No | 39 | [18 empty] | ppb
|  |  | PF*MS | [18 empty] | 1
1
| 1 | 1 | < 10 | [18 empty] | 1
| 1 | 2 | < 10 | [18 empty] | 26
| 1 | 3 | 14 | [18 empty] | 16
| 1 | 4 | 10 | [18 empty] | < 15
| 1 | 5 | 14 | [18 empty] | 13
| 1 | 6 | 12
| Mean |  | 12.5
| Median |  | 13
| Std Dev. |  | 1.9149
| Rel.Std.Dev. |  | 15.32%
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K in OREAS 45f (Certified Value 0.228 wt.%) | [20 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES | BF*OES/MS | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.22367 | 0.199 | 0.3 | 0.24863 | 0.2 | 0.24 | 0.31 | 0.2141 |  | 0.224 | 0.191 | 0.25 | 0.2329 | 0.246 | 0.2 | 0.241 | 0.2158 | 0.24055 |  | 1
| 1 | 2 | 0.21167 | 0.208 | 0.3 | 0.25776 | 0.2 | 0.26 | 0.29 | 0.2166 |  | 0.232 | 0.208 | 0.24 | 0.23668 | 0.245 | 0.2 | 0.232 | 0.2167 | 0.2361 |  | #N/A
| 1 | 3 | 0.21393 | 0.199 | 0.3 | 0.2532 | 0.2 | 0.22 | 0.33 | 0.2105 |  | 0.224 | 0.224 | 0.24 | 0.18772 | 0.257 | 0.2 | 0.249 | 0.2266 | 0.29056 |  | 16
| 1 | 4 | 0.21702 | 0.199 | 0.3 | 0.26357 | 0.2 | 0.25 | 0.3 | 0.2141 |  | 0.232 | 0.216 | 0.24 | 0.20437 | 0.245 | 0.3 | 0.241 | 0.225 | 0.2535 |  | 0.22828
| 1 | 5 | 0.21604 | 0.208 | 0.3 | 0.26606 | 0.2 | 0.25 | 0.29 | 0.2085 |  | 0.224 | 0.224 | 0.26 | 0.17428 | 0.253 | 0.2 | 0.257 | 0.2183 | 0.2268 |  | 52
| 1 | 6 | 0.2211 | 0.191 | 0.2 | 0.24822 | 0.2 | 0.23 | 0.31 | 0.2164 |  | 0.224 | 0.216 | 0.25 | 0.19261 | 0.24 | 0.3 | 0.232 | 0.2175 | 0.2299
| Mean |  | 0.21724 | 0.20067 | 0.28333 | 0.25624 | 0.2 | 0.24167 | 0.305 | 0.21337 |  | 0.22667 | 0.21317 | 0.24667 | 0.20476 | 0.24767 | 0.23333 | 0.242 | 0.21998 | 0.24624
| Median |  | 0.21653 | 0.199 | 0.3 | 0.25548 | 0.2 | 0.245 | 0.305 | 0.2141 |  | 0.224 | 0.216 | 0.245 | 0.19849 | 0.2455 | 0.2 | 0.241 | 0.2179 | 0.23833
| Std Dev. |  | 0.0044652 | 0.0064704 | 0.040825 | 0.0075359 | 3.0405e-17 | 0.01472 | 0.015166 | 0.0032445 |  | 0.0041312 | 0.0124 | 0.008165 | 0.025212 | 0.006186 | 0.05164 | 0.009757 | 0.0046093 | 0.023651
| Rel.Std.Dev. |  | 2.055% | 3.224% | 14.41% | 2.941% | 1.52e-14% | 6.091% | 4.972% | 1.521% |  | 1.823% | 5.817% | 3.31% | 12.31% | 2.498% | 22.13% | 4.032% | 2.095% | 9.605%
| PDM3 |  | -4.839% | -12.1% | 24.11% | 12.25% | -12.39% | 5.862% | 33.61% | -6.535% |  | -0.7084% | -6.622% | 8.053% | -10.31% | 8.491% | 2.212% | 6.008% | -3.636% | 7.864%
| Z-Score (Absolute) |  | 0.73 | 1.45 | NA | 0.94 | NA | 0.32 | 3.04 | 0.9 | NA | 0.32 | 0.91 | 0.53 | 1.27 | 0.58 | NA | 0.34 | 0.62 | 0.51
Results from laboratory 5, 12 & 33 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for La in OREAS 45f (Certified Value 15.7 ppm) | [20 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 05 | 10 | 14 | 17 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [8 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*OES | PF*MS | PF*MS | BF*MS | PF*OES | [8 empty] | 1
1
| 1 | 1 | 17.1 | 27.5031 | 15.65 |  | 16 | 17 | 11 | 17 | 13.5 | 16.3 | 13.923 | [8 empty] | 1
| 1 | 2 | 17.6 | 27.4684 | 14.55 |  | 16 | 16.8 | 15 | 16.7 | 15.2 | 18 | 11.467 | [8 empty] | #N/A
| 1 | 3 | 16.9 | 26.8634 | 14.75 |  | 17.5 | 16.2 | 15 | 17.9 | 14.7 | 18.6 | 20.302 | [8 empty] | 16
| 1 | 4 | 17.5 | 26.74 | 14.8 |  | 17 | 18.4 | 15 | 17.1 | 14.6 | 18 | 12.085 | [8 empty] | 15.6891
| 1 | 5 | 17.5 | 26.2226 | 14.25 |  | 17 | 15.9 | 14 | 17.1 | 13.4 | 15.7 | 13.31 | [8 empty] | 53
| 1 | 6 | 16.8 | 26.7518 | 15.1 |  | 16 | 15.8 | 14 | 16.7 | 13.7 | 16.3 | 13.393
| Mean |  | 17.2333 | 26.9249 | 14.85 |  | 16.5833 | 16.6833 | 14 | 17.0833 | 14.1833 | 17.15 | 14.08
| Median |  | 17.3 | 26.8076 | 14.775 |  | 16.5 | 16.5 | 14.5 | 17.05 | 14.15 | 17.15 | 13.3515
| Std Dev. |  | 0.34448 | 0.48802 | 0.4827 |  | 0.66458 | 0.96833 | 1.5492 | 0.44008 | 0.74677 | 1.1912 | 3.1815
| Rel.Std.Dev. |  | 1.999% | 1.813% | 3.251% |  | 4.008% | 5.804% | 11.07% | 2.576% | 5.265% | 6.946% | 22.6%
| PDM3 |  | 9.842% | 71.61% | -5.349% |  | 5.699% | 6.337% | -10.77% | 8.886% | -9.598% | 9.311% | -10.26%
| Z-Score (Absolute) |  | 0.34 | 5.82 | 1.01 | NA | 0.03 | 0.03 | 1.49 | 0.25 | 1.39 | 0.29 | 1.44
| Analytical results for Li in OREAS 45f (Certified Value 20.4 ppm) | [20 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 05 | 13 | 14 | 23 | 24 | 25 | 33 | 39 | [11 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [11 empty] | 1
2
| 1 | 1 | 20 | 20 | 21 | 18 | 20.8523 | 20.4 | 20.4 | 30 | [11 empty] | 1
| 1 | 2 | 20 | 22 | 22 | 17 | 21.2071 | 18.8 | 19.6 | 28 | [11 empty] | 21
| 1 | 3 | 30 | 18 | 23 | 20 | 19.7151 | 20.1 | 19 | 29 | [11 empty] | 16
| 1 | 4 | 20 | 21 | 23 | 31 | 20.0073 | 19.9 | 19.4 | 25 | [11 empty] | 20.4394
| 1 | 5 | 30 | 21 | 22 | 20 | 20.2488 | 20.5 | 20.4 | 32 | [11 empty] | 54
| 1 | 6 | 20 | 21 | 24 | 19 | 20.0865 | 20.1 | 21.3 | 32
| Mean |  | 23.3333 | 20.5 | 22.5 | 20.8333 | 20.3528 | 19.9667 | 20.0167 | 29.3333
| Median |  | 20 | 21 | 22.5 | 19.5 | 20.1676 | 20.1 | 20 | 29.5
| Std Dev. |  | 5.164 | 1.3784 | 1.0488 | 5.1153 | 0.56337 | 0.6121 | 0.84004 | 2.6583
| Rel.Std.Dev. |  | 22.13% | 6.724% | 4.661% | 24.55% | 2.768% | 3.066% | 4.197% | 9.062%
| PDM3 |  | 14.16% | 0.2967% | 10.08% | 1.928% | -0.4233% | -2.313% | -2.068% | 43.51%
| Z-Score (Absolute) |  | NA | 0 | 2.79 | 0.47 | 0.21 | 0.74 | 0.67 | 12.32
Results from laboratory 5 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 45f (Certified Value 0.26 ppm) | [20 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.25 | 0.13666 | 0.33 | 0.2 | 0.3 | 0.43 | 0.217 | < 0.25 | 0.25 | 0.211 | [9 empty] | 1
| 1 | 2 | 0.24 | 0.13977 | 0.22 | 0.2 | 0.29 | 0.4 | 0.226 | < 0.25 | 0.27 | 0.236 | [9 empty] | #N/A
| 1 | 3 | 0.24 | 0.13285 | 0.26 | 0.2 | 0.3 | 0.42 | 0.275 | < 0.25 | 0.31 | 0.214 | [9 empty] | 16
| 1 | 4 | 0.25 | 0.13404 | 0.2 | 0.2 | 0.29 | 0.44 | 0.277 | < 0.25 | 0.29 | 0.206 | [9 empty] | 0.25858
| 1 | 5 | 0.23 | 0.12889 | 0.23 | 0.2 | 0.29 | 0.41 | 0.291 | < 0.25 | 0.27 | 0.239 | [9 empty] | 55
| 1 | 6 | 0.26 | 0.1263 | 0.28 | 0.2 | 0.31 | 0.42 | 0.292 | < 0.25 | 0.27 | 0.195
| Mean |  | 0.245 | 0.13309 | 0.25333 | 0.2 | 0.29667 | 0.42 | 0.263 |  | 0.27667 | 0.21683
| Median |  | 0.245 | 0.13345 | 0.245 | 0.2 | 0.295 | 0.42 | 0.276 |  | 0.27 | 0.2125
| Std Dev. |  | 0.010488 | 0.0049404 | 0.047188 | 3.0405e-17 | 0.008165 | 0.014142 | 0.033015 |  | 0.020656 | 0.017291
| Rel.Std.Dev. |  | 4.281% | 3.712% | 18.63% | 1.52e-14% | 2.752% | 3.367% | 12.55% |  | 7.466% | 7.974%
| PDM3 |  | -5.253% | -48.53% | -2.03% | -22.66% | 14.73% | 62.42% | 1.708% |  | 6.993% | -16.15%
| Z-Score (Absolute) |  | 0.15 | 2.22 | 0 | NA | 0.8 | 3.08 | 0.18 | 2.37 | 0.43 | 0.67
Results from laboratory 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 45f (Certified Value 0.242 wt.%) | [20 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | BF*OES | BF*OES/MS | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.25142 | 0.217 | 0.25 | 0.24515 | 0.24 | 0.26 | 0.2659 |  | 0.24 | 0.247 | 0.259 | 0.23 | 0.24164 | 0.242 | 0.23 | 0.241 | 0.2424 | 0.2973 |  | 1
| 1 | 2 | 0.24798 | 0.235 | 0.25 | 0.24876 | 0.24 | 0.26 | 0.2689 |  | 0.23 | 0.241 | 0.271 | 0.23 | 0.23678 | 0.245 | 0.24 | 0.235 | 0.246 | 0.2902 |  | #N/A
| 1 | 3 | 0.25165 | 0.235 | 0.24 | 0.24816 | 0.24 | 0.24 | 0.2545 |  | 0.24 | 0.241 | 0.271 | 0.23 | 0.21296 | 0.243 | 0.23 | 0.241 | 0.2581 | 0.2926 |  | 16
| 1 | 4 | 0.24707 | 0.229 | 0.24 | 0.24967 | 0.25 | 0.25 | 0.2557 |  | 0.23 | 0.241 | 0.265 | 0.23 | 0.22443 | 0.238 | 0.24 | 0.235 | 0.2581 | 0.2977 |  | 0.24178
| 1 | 5 | 0.24966 | 0.229 | 0.24 | 0.26113 | 0.24 | 0.25 | 0.2492 |  | 0.25 | 0.241 | 0.271 | 0.23 | 0.21354 | 0.251 | 0.23 | 0.253 | 0.2454 | 0.2984 |  | 56
| 1 | 6 | 0.24751 | 0.229 | 0.24 | 0.24876 | 0.24 | 0.26 | 0.2611 |  | 0.25 | 0.241 | 0.265 | 0.22 | 0.21501 | 0.245 | 0.24 | 0.229 | 0.2491 | 0.2902
| Mean |  | 0.24922 | 0.229 | 0.24333 | 0.25027 | 0.24167 | 0.25333 | 0.25922 |  | 0.24 | 0.242 | 0.267 | 0.22833 | 0.22406 | 0.244 | 0.235 | 0.239 | 0.24985 | 0.2944
| Median |  | 0.24882 | 0.229 | 0.24 | 0.24876 | 0.24 | 0.255 | 0.2584 |  | 0.24 | 0.241 | 0.268 | 0.23 | 0.21972 | 0.244 | 0.235 | 0.238 | 0.24755 | 0.29495
| Std Dev. |  | 0.0020004 | 0.0065727 | 0.005164 | 0.0055412 | 0.0040825 | 0.008165 | 0.0074422 |  | 0.0089443 | 0.0024495 | 0.004899 | 0.0040825 | 0.012544 | 0.0042895 | 0.0054772 | 0.0081976 | 0.0067352 | 0.0038424
| Rel.Std.Dev. |  | 0.8027% | 2.87% | 2.122% | 2.214% | 1.689% | 3.223% | 2.871% |  | 3.727% | 1.012% | 1.835% | 1.788% | 5.598% | 1.758% | 2.331% | 3.43% | 2.696% | 1.305%
| PDM3 |  | 3.074% | -5.287% | 0.6414% | 3.511% | -0.04792% | 4.777% | 7.211% |  | -0.7372% | 0.08995% | 10.43% | -5.563% | -7.33% | 0.9171% | -2.805% | -1.151% | 3.337% | 21.76%
| Z-Score (Absolute) |  | 0.57 | 1.38 | 0 | 0.67 | 0.16 | 0.97 | 1.54 | NA | 0.32 | 0.11 | 2.33 | 1.46 | 1.87 | 0.06 | 0.81 | 0.4 | 0.64 | 4.96
| Analytical results for Mn in OREAS 45f (Certified Value 0.024 wt.%) | [20 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES | BF*OES | Various [a] | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES | BF*OES | PF*OES |  | 3
[13 empty] | see footer | [8 empty] | 3
| 1 | 1 | < 0.15 | 0.039 | 0.025 | 0.02362 | 0.024 | 0.02 | 0.025 |  | 0.025 | 0.023 | 0.023 | 0.0234 | 0.023917 | 0.0211 | 0.0282 | 0.023 | 0.0217 | 0.019673 |  | 1
| 1 | 2 | < 0.15 | 0.031 | 0.025 | 0.02362 | 0.024 | 0.02 | 0.024 |  | 0.024 | 0.023 | 0.023 | 0.0229 | 0.024199 | 0.0218 | 0.027 | 0.023 | 0.0225 | 0.018912 |  | #N/A
| 1 | 3 | < 0.15 | 0.031 | 0.025 | 0.02401 | 0.024 | 0.02 | 0.024 |  | 0.025 | 0.023 | 0.023 | 0.0219 | 0.022447 | 0.0211 | 0.0255 | 0.023 | 0.0232 | 0.019933 |  | 16
| 1 | 4 | < 0.15 | 0.031 | 0.026 | 0.0244 | 0.024 | 0.02 | 0.025 |  | 0.024 | 0.023 | 0.031 | 0.0234 | 0.023683 | 0.021 | 0.0255 | 0.023 | 0.0232 | 0.019704 |  | 0.023584
| 1 | 5 | < 0.15 | 0.031 | 0.025 | 0.02517 | 0.024 | 0.02 | 0.024 |  | 0.025 | 0.023 | 0.023 | 0.0227 | 0.022196 | 0.0221 | 0.0281 | 0.023 | 0.0217 | 0.019869 |  | 57
| 1 | 6 | < 0.15 | 0.039 | 0.025 | 0.02401 | 0.024 | 0.02 | 0.025 |  | 0.025 | 0.023 | 0.023 | 0.0242 | 0.023085 | 0.0215 | 0.026 | 0.023 | 0.0225 | 0.018981
| Mean |  |  | 0.033667 | 0.025167 | 0.024138 | 0.024 | 0.02 | 0.0245 |  | 0.024667 | 0.023 | 0.024333 | 0.023083 | 0.023254 | 0.021433 | 0.026717 | 0.023 | 0.022467 | 0.019512
| Median |  |  | 0.031 | 0.025 | 0.02401 | 0.024 | 0.02 | 0.0245 |  | 0.025 | 0.023 | 0.023 | 0.02315 | 0.023384 | 0.0213 | 0.0265 | 0.023 | 0.0225 | 0.019688
| Std Dev. |  |  | 0.0041312 | 0.00040825 | 0.00058362 | 3.8006e-18 | 0 | 0.00054772 |  | 0.0005164 | 3.8006e-18 | 0.003266 | 0.00077825 | 0.00081445 | 0.00044572 | 0.0012384 | 3.8006e-18 | 0.00067132 | 0.00044911
| Rel.Std.Dev. |  |  | 12.27% | 1.622% | 2.418% | 1.584e-14% | 0% | 2.236% |  | 2.094% | 1.652e-14% | 13.42% | 3.371% | 3.502% | 2.08% | 4.635% | 1.652e-14% | 2.988% | 2.302%
| PDM3 |  |  | 42.76% | 6.713% | 2.353% | 1.766% | -15.19% | 3.886% |  | 4.593% | -2.474% | 3.179% | -2.121% | -1.395% | -9.117% | 13.29% | -2.474% | -4.736% | -17.26%
| Z-Score (Absolute) |  | 30.89 | NA | 0.67 | 0.08 | 0 | 2.31 | 0.29 | NA | 0.39 | NA | NA | 0.53 | 0.43 | 1.48 | 1.57 | NA | 0.89 | 2.59
Results from laboratories 4, 20, 21 & 34 were removed due to their 0.01 wt.% reading resolution.
[a]: Lab 21: Rnd 01 - BF*MS, Rnd 01 - BF*OES
| Analytical results for Mo in OREAS 45f (Indicative Value 2.15 ppm) | [20 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 23 | 24 | 25 | 33 | 34 | 39 | [8 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [8 empty] | 2
2
| 1 | 1 | 2 | 2 | 2 |  | < 5 | 3 | 2.0038 | 2.4 | 3.4 | 3 | < 1 | [8 empty] | 1
| 1 | 2 | 2 | 2 | 2 |  | < 5 | 3 | 1.9934 | 3.7 | 3.8 | 3 | 1.384 | [8 empty] | 8
| 1 | 3 | 3 | 2 | 2 |  | < 5 | 3 | 1.7483 | 2.6 | 3.1 | 3 | 1 | [8 empty] | 16
| 1 | 4 | 3 | 2 | 2 |  | < 5 | 3 | 1.9485 | 2.4 | 2.8 | 3 | 1.135 | [8 empty] | 2.1527
| 1 | 5 | 3 | 3 | 3 |  | < 5 | 3 | 1.6502 | 2.6 | 3.8 | 3 | < 1 | [8 empty] | 14
| 1 | 6 | 3 | 2 | 3 |  | < 5 | 3 | 1.7986 | 2.5 | 2.9 | 3 | 1.203
| Mean |  | 2.6667 | 2.1667 | 2.3333 |  |  | 3 | 1.8571 | 2.7 | 3.3 | 3 | 1.1805
| Median |  | 3 | 2 | 2 |  |  | 3 | 1.8735 | 2.55 | 3.25 | 3 | 1.169
| Std Dev. |  | 0.5164 | 0.40825 | 0.5164 |  |  | 0 | 0.14596 | 0.498 | 0.43818 | 0 | 0.15976
| Rel.Std.Dev. |  | 19.36% | 18.84% | 22.13% |  |  | 0% | 7.859% | 18.44% | 13.28% | 0% | 13.53%
| PDM3 |  | 23.88% | 0.6489% | 8.391% |  |  | 39.36% | -13.73% | 25.42% | 53.3% | 39.36% | -45.16%
| Z-Score (Absolute) |  | NA | NA | NA | NA | 0 | NA | 0.67 | 0.21 | 0.84 | NA | 1.62
Results from laboratory 5, 13, 14, 23 & 34 were removed due to their 1 ppm or greater reading resolution.
| Analytical results for Na in OREAS 45f (Certified Value 0.07 wt.%) | [20 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 04 | 15 | 17 | 20 | 21 | 34 | 36 | [12 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES/MS | BF*OES | BF*MS | BF*OES | BF*OES | [12 empty] | 3
3
| 1 | 1 | 0.067 | 0.1859 |  | 0.067 | 0.074 | 0.074 | 0.0601 | [12 empty] | 1
| 1 | 2 | 0.067 | 0.1961 |  | 0.067 | 0.074 | 0.082 | 0.0601 | [12 empty] | #N/A
| 1 | 3 | 0.074 | 0.0842 |  | 0.059 | 0.067 | 0.074 | 0.0631 | [12 empty] | 16
| 1 | 4 | 0.074 | 0.117 |  | 0.059 | 0.074 | 0.067 | 0.0638 | [12 empty] | 0.069535
| 1 | 5 | 0.074 | 0.0942 |  | 0.067 | 0.074 | 0.111 | 0.0616 | [12 empty] | 58
| 1 | 6 | 0.082 | 0.1422 |  | 0.059 | 0.082 | 0.082 | 0.0623
| Mean |  | 0.073 | 0.1366 |  | 0.063 | 0.074167 | 0.081667 | 0.061833
| Median |  | 0.074 | 0.1296 |  | 0.063 | 0.074 | 0.078 | 0.06195
| Std Dev. |  | 0.0055857 | 0.046754 |  | 0.0043818 | 0.0047504 | 0.015449 | 0.0015332
| Rel.Std.Dev. |  | 7.652% | 34.23% |  | 6.955% | 6.405% | 18.92% | 2.48%
| PDM3 |  | 4.983% | 96.45% |  | -9.398% | 6.661% | 17.45% | -11.08%
| Z-Score (Absolute) |  | 0.04 | 4.58 | NA | 0.76 | 0.04 | 0.58 | 0.85
| Analytical results for Nb in OREAS 45f (Certified Value 23.7 ppm) | [20 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [5 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [5 empty] | 1
2
| 1 | 1 | 73 | 19 | 24 | 24.9 |  | 25 | 21.4 | 31.8 | 24 | 23.0844 | 25.4 | 25 | 24.67 | 26.171 | [5 empty] | 1
| 1 | 2 | 71 | 22 | 30 | 24.5 |  | 25 | 21.9 | 31.5 | 25 | 23.0913 | 25.3 | 25 | 24.68 | 26.651 | [5 empty] | 22
| 1 | 3 | 78 | 24 | 23 | 24.1 |  | 25 | 21.6 | 31.3 | 25 | 20.2926 | 25.6 | 25 | 25.55 | 27.458 | [5 empty] | 16
| 1 | 4 | 71 | 23 | 24 | 23.3 |  | 20 | 21.9 | 30.4 | 21 | 21.6411 | 25.4 | 25 | 24.69 | 24.958 | [5 empty] | 23.73
| 1 | 5 | 77 | 21 | 24 | 23.9 |  | 20 | 20.9 | 31.1 | 21 | 22.4198 | 25.9 | 25 | 24.64 | 28.709 | [5 empty] | 59
| 1 | 6 | 75 | 22 | 26 | 24.8 |  | 20 | 21.2 | 29.2 | 22 | 21.1624 | 25.9 | 30 | 24.61 | 24.599
| Mean |  | 74.1667 | 21.8333 | 25.1667 | 24.25 |  | 22.5 | 21.4833 | 30.8833 | 23 | 21.9486 | 25.5833 | 25.8333 | 24.8067 | 26.4243
| Median |  | 74 | 22 | 24 | 24.3 |  | 22.5 | 21.5 | 31.2 | 23 | 22.0304 | 25.5 | 25 | 24.675 | 26.411
| Std Dev. |  | 2.9944 | 1.7224 | 2.5626 | 0.60581 |  | 2.7386 | 0.39707 | 0.94956 | 1.8974 | 1.1201 | 0.26394 | 2.0412 | 0.36533 | 1.5418
| Rel.Std.Dev. |  | 4.037% | 7.889% | 10.18% | 2.498% |  | 12.17% | 1.848% | 3.075% | 8.249% | 5.103% | 1.032% | 7.902% | 1.473% | 5.835%
| PDM3 |  | 212.5% | -7.993% | 6.054% | 2.191% |  | -5.183% | -9.467% | 30.14% | -3.076% | -7.507% | 7.81% | 8.864% | 4.537% | 11.35%
| Z-Score (Absolute) |  | 18.42 | 1.11 | 0.13 | 0.21 | NA | 0.86 | 1.24 | 2.27 | 0.67 | 1.07 | 0.29 | 0.38 | 0 | 0.6
| Analytical results for Nd in OREAS 45f (Certified Value 13 ppm) | [20 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 1
2
| 1 | 1 | 13.1 | 21.3443 | 12.85 | 12.5 | 12 | 14.6 | 13.2 | < 10 | 13.8 | < 5 | [9 empty] | 1
| 1 | 2 | 13.2 | 21.5643 | 11.95 | 12 | 12.3 | 14.2 | 12.8 | < 10 | 14.2 | < 5 | [9 empty] | #N/A
| 1 | 3 | 12.6 | 21.0766 | 11.45 | 13 | 12.2 | 14.5 | 13.4 | < 10 | 14.3 | < 5 | [9 empty] | 16
| 1 | 4 | 12.8 | 20.877 | 11.85 | 12.5 | 12.6 | 14.3 | 12.8 | < 10 | 13.7 | < 5 | [9 empty] | 12.9512
| 1 | 5 | 12.7 | 19.416 | 11.8 | 13 | 11.9 | 14.5 | 13 | < 10 | 13.1 | < 5 | [9 empty] | 60
| 1 | 6 | 12.4 | 20.8381 | 12.55 | 12.5 | 11.6 | 14.9 | 12.6 | < 10 | 12.7 | < 5
| Mean |  | 12.8 | 20.8527 | 12.075 | 12.5833 | 12.1 | 14.5 | 12.9667 |  | 13.6333
| Median |  | 12.75 | 20.9768 | 11.9 | 12.5 | 12.1 | 14.5 | 12.9 |  | 13.75
| Std Dev. |  | 0.30332 | 0.7567 | 0.5213 | 0.37639 | 0.34641 | 0.24495 | 0.29439 |  | 0.62503
| Rel.Std.Dev. |  | 2.37% | 3.629% | 4.317% | 2.991% | 2.863% | 1.689% | 2.27% |  | 4.585%
| PDM3 |  | -1.167% | 61.01% | -6.765% | -2.84% | -6.572% | 11.96% | 0.1195% |  | 5.267%
| Z-Score (Absolute) |  | 0.09 | 7.06 | 0.53 | 0.09 | 0.51 | 1.56 | 0.24 | 6.66 | 0.81 | 8.82
| Analytical results for Ni in OREAS 45f (Certified Value 268 ppm) | [20 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 05 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES | BF*OES | BF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*MS | PF*OES |  |  |  |  | 0
0
| 1 | 1 | 264 | 231.8182 | 300 | 300 | 290 |  | 250 | 263 | 291 | 265 | 279.3936 | 248 | 260 | 280 | 223.815 |  |  |  |  | 1
| 1 | 2 | 266 | 245.4545 | 250 | 300 | 280 |  | 250 | 266 | 289 | 256 | 283.4295 | 301 | 315 | 273 | 208.393 |  |  |  |  | #N/A
| 1 | 3 | 273 | 236.3636 | 300 | 300 | 280 |  | 250 | 264 | 288 | 248 | 257.7324 | 257 | 280 | 283 | 255.783 |  |  |  |  | 16
| 1 | 4 | 266 | 254.5455 | 300 | 300 | 290 |  | 250 | 263 | 291 | 269 | 260.7501 | 254 | 275 | 273 | 238.426 |  |  |  |  | 268.4619
| 1 | 5 | 272 | 254.5455 | 300 | 300 | 280 |  | 250 | 257 | 280 | 262 | 270.1333 | 261 | 265 | 300 | 215.199 |  |  |  |  | 61
| 1 | 6 | 256 | 231.8182 | 300 | 300 | 290 |  | 250 | 261 | 290 | 265 | 283.5967 | 253 | 270 | 286 | 202
| Mean |  | 266.1667 | 242.4242 | 291.6667 | 300 | 285 |  | 250 | 262.3333 | 288.1667 | 260.8333 | 272.5059 | 262.3333 | 277.5 | 282.5 | 223.936
| Median |  | 266 | 240.9091 | 300 | 300 | 285 |  | 250 | 263 | 289.5 | 263.5 | 274.7635 | 255.5 | 272.5 | 281.5 | 219.507
| Std Dev. |  | 6.1455 | 10.6277 | 20.4124 | 0 | 5.4772 |  | 0 | 3.0768 | 4.1673 | 7.6267 | 11.418 | 19.4285 | 19.685 | 10.0548 | 20.1162
| Rel.Std.Dev. |  | 2.309% | 4.384% | 6.999% | 0% | 1.922% |  | 0% | 1.173% | 1.446% | 2.924% | 4.19% | 7.406% | 7.094% | 3.559% | 8.983%
| PDM3 |  | -0.8549% | -9.699% | 8.644% | 11.75% | 6.16% |  | -6.877% | -2.283% | 7.34% | -2.842% | 1.506% | -2.283% | 3.367% | 5.229% | -16.59%
| Z-Score (Absolute) |  | 0.15 | 1.26 | 1.04 | 1.43 | 0.73 | NA | 0.9 | 0.33 | 0.88 | 0.4 | 0.15 | 0.33 | 0.38 | 0.62 | 2.12
| Analytical results for P in OREAS 45f (Certified Value 0.03 wt.%) | [20 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 10 | 12 | 17 | 19 | 20 | 21 | 23 | 25 | 33 | 34 | 36 | 39 | [5 empty] | wt.%
|  |  | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | BF*OES/MS | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [5 empty] | 3
3
| 1 | 1 | 0.031 | 0.03 | 0.037336 | 0.04 |  | 0.03 | 0.026 | 0.031 | 0.03 | 0.0248 | 0.029 | N.A. | 0.0415 | 0.0304 | [5 empty] | 1
| 1 | 2 | 0.026 | 0.03 | 0.03559 | 0.04 |  | 0.03 | 0.022 | 0.026 | 0.03 | 0.0411 | 0.03 | N.A. | 0.0449 | 0.029757 | [5 empty] | #N/A
| 1 | 3 | 0.031 | 0.03 | 0.032969 | 0.04 |  | 0.04 | 0.022 | 0.026 | 0.03 | 0.0299 | 0.03 | N.A. | 0.0428 | 0.032156 | [5 empty] | 16
| 1 | 4 | 0.026 | 0.03 | 0.027948 | 0.04 |  | 0.03 | 0.026 | 0.026 | 0.03 | 0.0294 | 0.029 | N.A. | 0.0401 | 0.028041 | [5 empty] | 0.029817
| 1 | 5 | 0.026 | 0.03 | 0.032969 | 0.04 |  | 0.04 | 0.017 | 0.031 | 0.03 | 0.0313 | 0.029 | N.A. | 0.0449 | 0.028662 | [5 empty] | 62
| 1 | 6 | 0.026 | 0.04 | 0.040611 | 0.04 |  | 0.04 | 0.026 | 0.026 | 0.03 | 0.0297 | 0.032 | N.A. | 0.0345 | 0.028158
| Mean |  | 0.027667 | 0.031667 | 0.034571 | 0.04 |  | 0.035 | 0.023167 | 0.027667 | 0.03 | 0.031033 | 0.029833 |  | 0.04145 | 0.029529
| Median |  | 0.026 | 0.03 | 0.034279 | 0.04 |  | 0.035 | 0.024 | 0.026 | 0.03 | 0.0298 | 0.0295 |  | 0.04215 | 0.029209
| Std Dev. |  | 0.002582 | 0.0040825 | 0.0043398 | 0 |  | 0.0054772 | 0.0036009 | 0.002582 | 0 | 0.0054036 | 0.001169 |  | 0.0038924 | 0.0015847
| Rel.Std.Dev. |  | 9.332% | 12.89% | 12.55% | 0% |  | 15.65% | 15.54% | 9.332% | 0% | 17.41% | 3.919% |  | 9.391% | 5.367%
| PDM3 |  | -7.211% | 6.204% | 15.94% | 34.15% |  | 17.38% | -22.3% | -7.211% | 0.6147% | 4.08% | 0.05569% |  | 39.02% | -0.9652%
| Z-Score (Absolute) |  | 0.67 | 0.27 | 0.95 | 2.22 | NA | 1.05 | 1.7 | 0.67 | 0.12 | 0.12 | 0.16 | NA | 2.56 | 0.23
| Analytical results for Pb in OREAS 45f (Indicative Value 14.9 ppm) | [20 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 24 | 25 | 33 | 39 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | PF*OES | [10 empty] | 1
1
| 1 | 1 | 15 | < 20 | 14.3 |  | 10 | < 20 | 15 | < 20 | < 10 | [10 empty] | 1
| 1 | 2 | 16 | < 20 | 13.8 |  | 10 | < 20 | 15.1 | < 20 | < 10 | [10 empty] | 9
| 1 | 3 | 17 | < 20 | 13.4 |  | 10 | < 20 | 15.1 | < 20 | < 10 | [10 empty] | 16
| 1 | 4 | 15 | < 20 | 13.7 |  | 10 | < 20 | 14.9 | < 20 | < 10 | [10 empty] | 14.8667
| 1 | 5 | 19 | < 20 | 13.4 |  | 10 | < 20 | 14.8 | < 20 | < 10 | [10 empty] | 15
| 1 | 6 | 14 | < 20 | 13.3 |  | 10 | < 20 | 14.8 | < 20 | < 10
| Mean |  | 16 |  | 13.65 |  | 10 |  | 14.95
| Median |  | 15.5 |  | 13.55 |  | 10 |  | 14.95
| Std Dev. |  | 1.7889 |  | 0.37283 |  | 0 |  | 0.13784
| Rel.Std.Dev. |  | 11.18% |  | 2.731% |  | 0% |  | 0.922%
| PDM3 |  | 7.623% |  | -8.184% |  | -32.74% |  | 0.5605%
| Z-Score (Absolute) |  | 1.11 | 0 | 0.67 | NA | NA | 0 | 0.91 | 0 | 0.92
Results from laboratory 19 were removed due to their 10 ppm reading resolution.
| Analytical results for Pr in OREAS 45f (Certified Value 3.44 ppm) | [20 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 2
3
| 1 | 1 | 3.34 | 5.7017 | 3.24 | 3.6 | 3.17 | 3.26 | 3.67 | < 10 | 3.63 | < 1 | [9 empty] | 1
| 1 | 2 | 3.44 | 5.6959 | 3.1 | 3.6 | 3.35 | 3.49 | 3.64 | < 10 | 3.67 | < 1 | [9 empty] | #N/A
| 1 | 3 | 3.26 | 5.5404 | 2.94 | 3.8 | 3.25 | 3.5 | 3.93 | < 10 | 3.81 | < 1 | [9 empty] | 16
| 1 | 4 | 3.49 | 5.5522 | 3.06 | 3.6 | 3.27 | 3.52 | 3.71 | < 10 | 3.53 | < 1 | [9 empty] | 3.4421
| 1 | 5 | 3.37 | 5.06 | 3.46 | 3.6 | 3.09 | 3.46 | 3.7 | < 10 | 3.48 | < 1 | [9 empty] | 63
| 1 | 6 | 3.2 | 5.5078 | 3.35 | 3.4 | 3.18 | 3.45 | 3.57 | < 10 | 3.39 | < 1
| Mean |  | 3.35 | 5.5097 | 3.1917 | 3.6 | 3.2183 | 3.4467 | 3.7033 |  | 3.585
| Median |  | 3.355 | 5.5463 | 3.17 | 3.6 | 3.215 | 3.475 | 3.685 |  | 3.58
| Std Dev. |  | 0.10844 | 0.23517 | 0.19416 | 0.12649 | 0.090866 | 0.095009 | 0.12193 |  | 0.14963
| Rel.Std.Dev. |  | 3.237% | 4.268% | 6.083% | 3.514% | 2.823% | 2.757% | 3.292% |  | 4.174%
| PDM3 |  | -2.677% | 60.06% | -7.277% | 4.586% | -6.502% | 0.1314% | 7.588% |  | 4.15%
| Z-Score (Absolute) |  | 0.46 | 5.54 | 0.9 | 0.23 | 0.83 | 0.19 | 0.52 | 4.13 | 0.19 | 8.39
| Analytical results for Pt in OREAS 45f (Indicative Value < 50 ppb) | [20 empty] | Indicative
Pt | Round | Replicate | Lab | [18 empty] | 1
| No | No | 39 | [18 empty] | ppb
|  |  | PF*MS | [18 empty] | 1
1
| 1 | 1 | < 50 | [18 empty] | 1
| 1 | 2 | < 50 | [18 empty] | 10
| 1 | 3 | < 50 | [18 empty] | 16
| 1 | 4 | < 50 | [18 empty] | < 50
| 1 | 5 | < 50 | [18 empty] | 16
| 1 | 6 | < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 45f (Certified Value 31.6 ppm) | [20 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [6 empty] | 1
2
| 1 | 1 | 31.4 | 31.5 | 32.6 |  | 30.5 | 27.3 | 36.3 | 29 | 31.7099 | 35.9 | 35 | 31.7 | 31.046 | [6 empty] | 1
| 1 | 2 | 31.7 | 31.7 | 33.4 |  | 30.5 | 29 | 37.5 | 31 | 32.334 | 34.1 | 34 | 31.2 | 30.027 | [6 empty] | 2
| 1 | 3 | 31.7 | 30.9 | 33.8 |  | 31 | 28 | 35.6 | 30.2 | 31.3137 | 35.3 | 32.2 | 31.6 | 29.086 | [6 empty] | 16
| 1 | 4 | 32.1 | 31.8 | 31.1 |  | 31 | 28.6 | 34.6 | 28.4 | 31.5963 | 35.1 | 33.1 | 32.1 | 29.419 | [6 empty] | 31.5532
| 1 | 5 | 31.7 | 32.7 | 32 |  | 32 | 26.8 | 37.8 | 30.4 | 32.1588 | 35.1 | 34.1 | 31.8 | 32.977 | [6 empty] | 64
| 1 | 6 | 31.8 | 32.6 | 33.3 |  | 30 | 27.3 | 34.7 | 30.1 | 31.9586 | 35.4 | 32.8 | 31.2 | 28.283
| Mean |  | 31.7333 | 31.8667 | 32.7 |  | 30.8333 | 27.8333 | 36.0833 | 29.85 | 31.8452 | 35.15 | 33.5333 | 31.6 | 30.1397
| Median |  | 31.7 | 31.75 | 32.95 |  | 30.75 | 27.65 | 35.95 | 30.15 | 31.8342 | 35.2 | 33.55 | 31.65 | 29.723
| Std Dev. |  | 0.22509 | 0.68313 | 1.0119 |  | 0.68313 | 0.8501 | 1.3674 | 0.96281 | 0.3776 | 0.59245 | 1.0191 | 0.35214 | 1.671
| Rel.Std.Dev. |  | 0.7093% | 2.144% | 3.095% |  | 2.216% | 3.054% | 3.789% | 3.225% | 1.186% | 1.685% | 3.039% | 1.114% | 5.544%
| PDM3 |  | 0.571% | 0.9936% | 3.635% |  | -2.281% | -11.79% | 14.36% | -5.398% | 0.9255% | 11.4% | 6.276% | 0.1484% | -4.48%
| Z-Score (Absolute) |  | 0.03 | 0.04 | 0.47 | NA | 0.49 | 2.05 | 2.22 | 1 | 0.03 | 1.74 | 0.9 | 0.1 | 0.85
| Analytical results for Re in OREAS 45f (Indicative Value < 0.01 ppm) | [20 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 13 | 14 | 19 | 24 | 39 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [14 empty] | 3
3
| 1 | 1 | < 0.1 | < 0.01 | < 1 | < 0.1 | < 0.01 | [14 empty] | 1
| 1 | 2 | < 0.1 | < 0.01 | < 1 | < 0.1 | < 0.01 | [14 empty] | 11
| 1 | 3 | < 0.1 | < 0.01 | < 1 | < 0.1 | < 0.01 | [14 empty] | 16
| 1 | 4 | < 0.1 | < 0.01 | < 1 | < 0.1 | < 0.01 | [14 empty] | < 0.01
| 1 | 5 | < 0.1 | < 0.01 | < 1 | < 0.1 | < 0.01 | [14 empty] | 17
| 1 | 6 | < 0.1 | < 0.01 | < 1 | < 0.1 | < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA
| Analytical results for Rh in OREAS 45f (Indicative Value < 30 ppb) | [20 empty] | Indicative
Rh | Round | Replicate | Lab | [18 empty] | 1
| No | No | 39 | [18 empty] | ppb
|  |  | PF*MS | [18 empty] | 1
1
| 1 | 1 | < 30 | [18 empty] | 1
| 1 | 2 | < 30 | [18 empty] | 12
| 1 | 3 | < 30 | [18 empty] | 16
| 1 | 4 | < 30 | [18 empty] | < 30
| 1 | 5 | < 30 | [18 empty] | 18
| 1 | 6 | < 30
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ru in OREAS 45f (Indicative Value < 30 ppb) | [20 empty] | Indicative
Ru | Round | Replicate | Lab | [18 empty] | 1
| No | No | 39 | [18 empty] | ppb
|  |  | PF*MS | [18 empty] | 1
1
| 1 | 1 | < 30 | [18 empty] | 1
| 1 | 2 | < 30 | [18 empty] | 13
| 1 | 3 | < 30 | [18 empty] | 16
| 1 | 4 | < 30 | [18 empty] | < 30
| 1 | 5 | < 30 | [18 empty] | 19
| 1 | 6 | < 30
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 45f (Indicative Value 0.034 wt.%) | [20 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 12 | 13 | 17 | 23 | 24 | 33 | 39 | [12 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | [12 empty] | 3
3
| 1 | 1 | 0.04 | < 0.05 |  | 0.06 | < 0.05 | 0.03 | 0.0226 | [12 empty] | 1
| 1 | 2 | 0.04 | < 0.05 |  | 0.07 | < 0.05 | 0.03 | 0.0239 | [12 empty] | 28
| 1 | 3 | 0.04 | < 0.05 |  | < 0.05 | < 0.05 | 0.03 | 0.020135 | [12 empty] | 16
| 1 | 4 | 0.04 | < 0.05 |  | 0.06 | < 0.05 | 0.03 | 0.0233 | [12 empty] | 0.034351
| 1 | 5 | 0.04 | < 0.05 |  | < 0.05 | < 0.05 | 0.03 | 0.0263 | [12 empty] | 13
| 1 | 6 | 0.04 | < 0.05 |  | < 0.05 | < 0.05 | 0.03 | 0.0232
| Mean |  | 0.04 |  |  | 0.063333 |  | 0.03 | 0.023239
| Median |  | 0.04 |  |  | 0.06 |  | 0.03 | 0.02325
| Std Dev. |  | 0 |  |  | 0.0057735 |  | 0 | 0.0019926
| Rel.Std.Dev. |  | 0% |  |  | 9.116% |  | 0% | 8.574%
| PDM3 |  | 16.44% |  |  | 84.37% |  | -12.67% | -32.35%
| Z-Score (Absolute) |  | 2.49 | 0.5 | NA | 3.32 | 0.5 | 0.5 | 0.85
| Analytical results for Sb in OREAS 45f (Certified Value 0.63 ppm) | [20 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 23 | 24 | 25 | 33 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.4 | 0.8 | 0.5 |  | < 10 | 0.8 | 0.67234 | 0.736 | 0.5 | 0.3 | [9 empty] | 1
| 1 | 2 | 0.5 | 0.6 | 0.6 |  | < 10 | 0.5 | 0.84891 | 0.641 | 0.6 | 0.5 | [9 empty] | 26
| 1 | 3 | 0.5 | 0.8 | 0.6 |  | < 10 | 0.7 | 0.79352 | 0.672 | 0.5 | 0.2 | [9 empty] | 16
| 1 | 4 | 0.5 | 0.7 | 0.9 |  | < 10 | 0.8 | 0.79919 | 0.803 | 0.5 | 0.504 | [9 empty] | 0.62922
| 1 | 5 | 0.5 | 0.7 | 0.8 |  | < 10 | < 0.5 | 0.65929 | 0.68 | 0.5 | 0.5 | [9 empty] | 65
| 1 | 6 | 0.6 | 0.8 | 0.4 |  | < 10 | 0.7 | 0.75038 | 0.782 | 0.6 | 0.5
| Mean |  | 0.5 | 0.73333 | 0.63333 |  |  | 0.7 | 0.75394 | 0.719 | 0.53333 | 0.41733
| Median |  | 0.5 | 0.75 | 0.6 |  |  | 0.7 | 0.77195 | 0.708 | 0.5 | 0.5
| Std Dev. |  | 0.063246 | 0.08165 | 0.18619 |  |  | 0.12247 | 0.075187 | 0.065005 | 0.05164 | 0.13343
| Rel.Std.Dev. |  | 12.65% | 11.13% | 29.4% |  |  | 17.5% | 9.973% | 9.041% | 9.682% | 31.97%
| PDM3 |  | -20.54% | 16.55% | 0.6542% |  |  | 11.25% | 19.82% | 14.27% | -15.24% | -33.67%
| Z-Score (Absolute) |  | 0.9 | 0.67 | 0 | NA | 29.44 | 0.06 | 0.81 | 0.58 | 0.67 | 1.46
| Analytical results for Sc in OREAS 45f (Certified Value 35.8 ppm) | [20 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 12 | 17 | 19 | 20 | 21 | 24 | 25 | 33 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES | BF*OES/MS | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 1
2
| 1 | 1 | 40 | 40 |  | 30 | 36 | 35 | 25.1646 | 35.5 | 33 | 34.213 | [9 empty] | 1
| 1 | 2 | 39 | 40 |  | 30 | 36 | 35 | 29.9673 | 36.3 | 34 | 37.931 | [9 empty] | #N/A
| 1 | 3 | 39 | 40 |  | 30 | 36 | 34 | 28.8372 | 34.9 | 33 | 36 | [9 empty] | 16
| 1 | 4 | 39 | 40 |  | 30 | 36 | 33 | 27.4193 | 35.8 | 35 | 39 | [9 empty] | 35.7957
| 1 | 5 | 39 | 40 |  | 40 | 36 | 34 | 26.2803 | 36.7 | 33 | 36 | [9 empty] | 66
| 1 | 6 | 38 | 40 |  | 30 | 36 | 33 | 27.4193 | 36.3 | 33 | 35
| Mean |  | 39 | 40 |  | 31.6667 | 36 | 34 | 27.5147 | 35.9167 | 33.5 | 36.3573
| Median |  | 39 | 40 |  | 30 | 36 | 34 | 27.4193 | 36.05 | 33 | 36
| Std Dev. |  | 0.63246 | 0 |  | 4.0825 | 0 | 0.89443 | 1.7221 | 0.65243 | 0.83666 | 1.7976
| Rel.Std.Dev. |  | 1.622% | 0% |  | 12.89% | 0% | 2.631% | 6.259% | 1.817% | 2.497% | 4.944%
| PDM3 |  | 8.952% | 11.75% |  | -11.53% | 0.5708% | -5.016% | -23.13% | 0.338% | -6.413% | 1.569%
| Z-Score (Absolute) |  | 1.08 | NA | NA | NA | 0.03 | 0.67 | 2.96 | 0 | 0.85 | 0.16
Results from laboratories 12 & 19 were removed due to their 10 ppm reading resolution.
| Analytical results for Se in OREAS 45f (Indicative Value < 20 ppm) | [20 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 13 | 14 | 24 | 33 | 39 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [14 empty] | 1
1
| 1 | 1 | < 20 | 4 | < 20 | < 50 | 1.799 | [14 empty] | 1
| 1 | 2 | < 20 | < 3 | < 20 | < 50 | 1.788 | [14 empty] | 8
| 1 | 3 | < 20 | < 3 | < 20 | < 50 | 1.423 | [14 empty] | 16
| 1 | 4 | < 20 | < 3 | < 20 | < 50 | 1.628 | [14 empty] | < 20
| 1 | 5 | < 20 | < 3 | < 20 | < 50 | 1.699 | [14 empty] | 14
| 1 | 6 | < 20 | 9 | < 20 | < 50 | 1.65
| Mean |  |  | 6.5 |  |  | 1.6645
| Median |  |  | 6.5 |  |  | 1.6745
| Std Dev. |  |  | 3.5355 |  |  | 0.1374
| Rel.Std.Dev. |  |  | 54.39% |  |  | 8.255%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 0 | 1.48 | 0.82
| Analytical results for Si in OREAS 45f (Certified Value 19.72 wt.%) | [20 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 15 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 19.7883 | 18.418 | 20.4 | 19.2997 | 20.4 | 19.7 | 18.75 | 20.2 | 19.643 | 18.212 | 19.3 | 19.4446 | 19.7 | 20.4 | 20.695 | 19.3063 | 19.8497 |  |  | 1
| 1 | 2 | 19.7665 | 18.329 | 20.6 | 19.219 | 20.5 | 19.9 | 18.81 | 19.6 | 19.605 | 18.427 | 19.1 | 19.6454 | 19.6 | 20.6 | 20.115 | 19.2595 | 20.0674 |  |  | #N/A
| 1 | 3 | 19.7669 | 18.722 | 20.3 | 19.1585 | 20.2 | 19.1 | 19.74 | 20.3 | 19.587 | 18.217 | 18.9 | 18.939 | 20 | 20.1 | 20.554 | 20.4282 | 19.3705 |  |  | 16
| 1 | 4 | 19.7988 | 18.311 | 20.6 | 19.1585 | 20.3 | 19.5 | 19.58 | 19.8 | 19.554 | 18.002 | 19.8 | 20.0728 | 19.6 | 20.6 | 20.381 | 20.4282 | 20.0196 |  |  | 19.7176
| 1 | 5 | 19.7778 | 18.661 | 20.4 | 19.3076 | 20.5 | 20.2 | 19.05 | 19.4 | 19.559 | 18.184 | 19.4 | 20.0221 | 20.1 | 20.2 | 20.148 | 19.5868 | 19.4957 |  |  | 67
| 1 | 6 | 19.8199 | 18.839 | 20.6 | 19.3579 | 20.5 | 19.4 | 19.58 | 20.3 | 19.648 | 18.311 | 19.3 | 18.4595 | 19.7 | 20.6 | 20.265 | 19.54 | 18.4923
| Mean |  | 19.7864 | 18.5467 | 20.4833 | 19.2502 | 20.4 | 19.6333 | 19.2517 | 19.9333 | 19.5993 | 18.2255 | 19.3 | 19.4306 | 19.7833 | 20.4167 | 20.3597 | 19.7582 | 19.5492
| Median |  | 19.7831 | 18.5395 | 20.5 | 19.2594 | 20.45 | 19.6 | 19.315 | 20 | 19.596 | 18.2145 | 19.3 | 19.545 | 19.7 | 20.5 | 20.323 | 19.5634 | 19.6727
| Std Dev. |  | 0.020621 | 0.22304 | 0.13292 | 0.083818 | 0.12649 | 0.38816 | 0.43393 | 0.38816 | 0.040352 | 0.14131 | 0.30332 | 0.6314 | 0.2137 | 0.22286 | 0.22991 | 0.53436 | 0.58855
| Rel.Std.Dev. |  | 0.1042% | 1.203% | 0.6489% | 0.4354% | 0.6201% | 1.977% | 2.254% | 1.947% | 0.2059% | 0.7753% | 1.572% | 3.25% | 1.08% | 1.092% | 1.129% | 2.705% | 3.011%
| PDM3 |  | 0.3488% | -5.939% | 3.883% | -2.37% | 3.461% | -0.4274% | -2.363% | 1.094% | -0.5998% | -7.567% | -2.118% | -1.456% | 0.3333% | 3.545% | 3.256% | 0.2057% | -0.8542%
| Z-Score (Absolute) |  | 0.31 | 2.2 | 1.72 | 0.78 | 1.55 | 0 | 0.77 | 0.61 | 0.07 | 2.85 | 0.67 | 0.41 | 0.3 | 1.58 | 1.47 | 0.25 | 0.17
| Analytical results for Sm in OREAS 45f (Certified Value 2.6 ppm) | [20 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 2
3
| 1 | 1 | 2.6 | 4.1391 | 2.45 | 2.5 | 2.43 | 2.71 | 2.75 | < 100 | 2.6 | < 1 | [9 empty] | 1
| 1 | 2 | 2.9 | 4.2007 | 2.25 | 2.5 | 2.47 | 2.7 | 2.59 | < 100 | 2.8 | < 1 | [9 empty] | #N/A
| 1 | 3 | 2.7 | 4.0664 | 2.35 | 2.5 | 2.51 | 2.68 | 2.83 | < 100 | 2.8 | < 1 | [9 empty] | 16
| 1 | 4 | 2.7 | 4.0332 | 2.39 | 2.5 | 2.5 | 2.78 | 2.69 | < 100 | 2.9 | < 1 | [9 empty] | 2.6018
| 1 | 5 | 2.9 | 3.93 | 2.48 | 2.5 | 2.34 | 2.68 | 2.78 | < 100 | 2.6 | < 1 | [9 empty] | 68
| 1 | 6 | 2.5 | 4.0194 | 2.55 | 2.5 | 2.51 | 2.88 | 2.6 | < 100 | 2.4 | < 1
| Mean |  | 2.7167 | 4.0648 | 2.4117 | 2.5 | 2.46 | 2.7383 | 2.7067 |  | 2.6833
| Median |  | 2.7 | 4.0498 | 2.42 | 2.5 | 2.485 | 2.705 | 2.72 |  | 2.7
| Std Dev. |  | 0.16021 | 0.095085 | 0.10553 | 0 | 0.066332 | 0.078592 | 0.097707 |  | 0.18348
| Rel.Std.Dev. |  | 5.897% | 2.339% | 4.376% | 0% | 2.696% | 2.87% | 3.61% |  | 6.838%
| PDM3 |  | 4.416% | 56.23% | -7.306% | -3.911% | -5.449% | 5.249% | 4.032% |  | 3.135%
| Z-Score (Absolute) |  | 0.07 | 4.3 | 0.89 | 0.61 | 0.74 | 0.14 | 0.04 | 148.36 | 0.04 | 6.88
| Analytical results for Sn in OREAS 45f (Certified Value 3.37 ppm) | [20 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 24 | 25 | 33 | 34 | 39 | [7 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [7 empty] | 2
2
| 1 | 1 | 3 | < 100 | < 3 |  | < 10 | 3 | 4 | < 100 | 2.63 | < 100 | 3.5 | 3.643 | [7 empty] | 1
| 1 | 2 | 3 | < 100 | 3 |  | < 10 | 3 | 4 | < 100 | 4.41 | < 100 | 3.3 | 3.676 | [7 empty] | 3
| 1 | 3 | 3 | < 100 | < 3 |  | < 10 | 3 | 3 | < 100 | 5.03 | < 100 | 3.7 | 3.59 | [7 empty] | 16
| 1 | 4 | 3 | < 100 | < 3 |  | < 10 | 2 | 4 | < 100 | 4.47 | < 100 | 2.9 | 3.69 | [7 empty] | 3.369
| 1 | 5 | 3 | < 100 | < 3 |  | < 10 | 3 | 4 | < 100 | 4.27 | < 100 | 3.3 | 3.32 | [7 empty] | 69
| 1 | 6 | 2 | < 100 | < 3 |  | < 10 | 2 | 4 | < 100 | 3.69 | < 100 | 3.2 | 3.434
| Mean |  | 2.8333 |  | 3 |  |  | 2.6667 | 3.8333 |  | 4.0833 |  | 3.3167 | 3.5588
| Median |  | 3 |  | 3 |  |  | 3 | 4 |  | 4.34 |  | 3.3 | 3.6165
| Std Dev. |  | 0.40825 |  |  |  |  | 0.5164 | 0.40825 |  | 0.83116 |  | 0.27142 | 0.14941
| Rel.Std.Dev. |  | 14.41% |  |  |  |  | 19.36% | 10.65% |  | 20.35% |  | 8.183% | 4.198%
| PDM3 |  | -15.9% |  | -10.95% |  |  | -20.85% | 13.78% |  | 21.2% |  | -1.553% | 5.635%
| Z-Score (Absolute) |  | 0.58 | 26.68 | 1.2 | NA | 0.67 | 0.67 | 0 | 26.68 | 0.14 | 26.68 | 0.3 | 0.16
| Analytical results for Sr in OREAS 45f (Certified Value 26.9 ppm) | [20 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 24 | 25 | 33 | 34 | 36 | 39 | [5 empty] | ppm
|  |  | BF*OES | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [5 empty] | 1
1
| 1 | 1 | 34 | 30 | 24 | 20 |  | 31 | 23.8 | 22 | 25.7572 | 24.3 | 30 | < 10 | < 42 | 30 | [5 empty] | 1
| 1 | 2 | 35 | 30 | 24 | 20 |  | 30 | 24.8 | 22 | 27.49 | 26.1 | 30 | < 10 | < 42 | 29 | [5 empty] | #N/A
| 1 | 3 | 37 | 30 | 24 | 20 |  | 31 | 24.8 | 22 | 24.4561 | 27.4 | 30 | 10 | < 42 | 28 | [5 empty] | 16
| 1 | 4 | 36 | 30 | 24 | 30 |  | 30 | 25.8 | 22 | 25.9193 | 25.4 | 30 | < 10 | < 42 | 29 | [5 empty] | 26.8555
| 1 | 5 | 35 | 30 | 25 | 30 |  | 32 | 23.6 | 22 | 24.5689 | 26.1 | 30 | 21 | < 42 | 28 | [5 empty] | 70
| 1 | 6 | 38 | 30 | 25 | 30 |  | 35 | 22.1 | 22 | 27.6401 | 25.5 | 30 | < 10 | < 42 | 28
| Mean |  | 35.8333 | 30 | 24.3333 | 25 |  | 31.5 | 24.15 | 22 | 25.9719 | 25.8 | 30 | 15.5 |  | 28.6667
| Median |  | 35.5 | 30 | 24 | 25 |  | 31 | 24.3 | 22 | 25.8382 | 25.8 | 30 | 15.5 |  | 28.5
| Std Dev. |  | 1.472 | 0 | 0.5164 | 5.4772 |  | 1.8708 | 1.2802 | 0 | 1.3713 | 1.0237 | 0 | 7.7782 |  | 0.8165
| Rel.Std.Dev. |  | 4.108% | 0% | 2.122% | 21.91% |  | 5.939% | 5.301% | 0% | 5.28% | 3.968% | 0% | 50.18% |  | 2.848%
| PDM3 |  | 33.43% | 11.71% | -9.392% | -6.909% |  | 17.29% | -10.07% | -18.08% | -3.29% | -3.93% | 11.71% | -42.28% |  | 6.744%
| Z-Score (Absolute) |  | 1.78 | NA | 0.26 | NA | NA | 1.01 | 0.29 | 0.67 | 0.03 | 0 | 0.75 | 3.07 | 0.83 | 0.51
Results from laboratories 5 & 14 were removed due to their 10 ppm reading resolution.
| Analytical results for Ta in OREAS 45f (Certified Value 1.56 ppm) | [20 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 1.3 | 1.6 | 1.56 |  | 1.5 | 1.5 | 1.9 | 1.6 | 1.9244 | 1.58 | < 10 | 1.66 | 6.865 | [6 empty] | 1
| 1 | 2 | 1.4 | 1.6 | 1.34 |  | 1.5 | 1.4 | 1.8 | 2.9 | 2.0807 | 1.55 | < 10 | 1.64 | 6.611 | [6 empty] | 22
| 1 | 3 | 1.4 | 1.5 | 1.39 |  | 1.5 | 1.3 | 1.7 | 3.9 | 1.4365 | 1.64 | < 10 | 1.67 | 7.69 | [6 empty] | 16
| 1 | 4 | 1.4 | 1.7 | 1.47 |  | 1.5 | 1.5 | 1.6 | 2.2 | 1.7886 | 1.57 | < 10 | 1.61 | 7.491 | [6 empty] | 1.5551
| 1 | 5 | 1.3 | 1.5 | 1.43 |  | 1.5 | 1.6 | 1.8 | 2.3 | 1.6602 | 1.58 | < 10 | 1.68 | 6.98 | [6 empty] | 71
| 1 | 6 | 1.3 | 1.6 | 1.56 |  | 1 | 1.5 | 1.6 | 2.5 | 1.5214 | 1.61 | < 10 | 1.64 | 7.254
| Mean |  | 1.35 | 1.5833 | 1.4583 |  | 1.4167 | 1.4667 | 1.7333 | 2.5667 | 1.7353 | 1.5883 |  | 1.65 | 7.1485
| Median |  | 1.35 | 1.6 | 1.45 |  | 1.5 | 1.5 | 1.75 | 2.4 | 1.7244 | 1.58 |  | 1.65 | 7.117
| Std Dev. |  | 0.054772 | 0.075277 | 0.089759 |  | 0.20412 | 0.10328 | 0.12111 | 0.77889 | 0.24438 | 0.031885 |  | 0.025298 | 0.4048
| Rel.Std.Dev. |  | 4.057% | 4.754% | 6.155% |  | 14.41% | 7.042% | 6.987% | 30.35% | 14.08% | 2.007% |  | 1.533% | 5.663%
| PDM3 |  | -13.19% | 1.813% | -6.225% |  | -8.904% | -5.689% | 11.46% | 65.04% | 11.59% | 2.135% |  | 6.1% | 359.7%
| Z-Score (Absolute) |  | 1.16 | 0.15 | 0.69 | NA | 0.87 | 0.66 | 0.49 | 4.08 | 0.5 | 0.13 | 14.55 | 0.13 | 23.8
| Analytical results for Tb in OREAS 45f (Certified Value 0.41 ppm) | [20 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | [9 empty] | 2
3
| 1 | 1 | 0.39 | 0.37644 | 0.4 | 0.4 | 0.45 | 0.39 | 0.422 | 0.4 | 0.44 | < 5 | [9 empty] | 1
| 1 | 2 | 0.39 | 0.3818 | 0.44 | 0.4 | 0.44 | 0.4 | 0.435 | 0.4 | 0.47 | < 5 | [9 empty] | #N/A
| 1 | 3 | 0.4 | 0.36929 | 0.37 | 0.4 | 0.44 | 0.42 | 0.431 | 0.3 | 0.5 | < 5 | [9 empty] | 16
| 1 | 4 | 0.41 | 0.36314 | 0.43 | 0.4 | 0.47 | 0.41 | 0.421 | 0.3 | 0.41 | < 5 | [9 empty] | 0.4094
| 1 | 5 | 0.43 | 0.35065 | 0.34 | 0.4 | 0.45 | 0.38 | 0.442 | 0.4 | 0.44 | < 5 | [9 empty] | 72
| 1 | 6 | 0.42 | 0.351 | 0.45 | 0.4 | 0.45 | 0.41 | 0.428 | 0.4 | 0.44 | < 5
| Mean |  | 0.40667 | 0.36539 | 0.405 | 0.4 | 0.45 | 0.40167 | 0.42983 | 0.36667 | 0.45
| Median |  | 0.405 | 0.36622 | 0.415 | 0.4 | 0.45 | 0.405 | 0.4295 | 0.4 | 0.44
| Std Dev. |  | 0.01633 | 0.01293 | 0.043243 | 6.0809e-17 | 0.010954 | 0.01472 | 0.0079854 | 0.05164 | 0.030984
| Rel.Std.Dev. |  | 4.016% | 3.539% | 10.68% | 1.52e-14% | 2.434% | 3.665% | 1.858% | 14.08% | 6.885%
| PDM3 |  | -0.6683% | -10.75% | -1.075% | -2.297% | 9.916% | -1.89% | 4.99% | -10.44% | 9.916%
| Z-Score (Absolute) |  | 0 | 0.7 | 0.03 | NA | 0.73 | 0.08 | 0.39 | 0.67 | 0.73 | 35.29
Results from laboratory 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 45f (Indicative Value < 1 ppm) | [20 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 13 | 14 | 24 | 33 | 39 | [14 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | [14 empty] | 2
2
| 1 | 1 | < 1 | < 0.5 | < 2 | < 1 | < 0.5 | [14 empty] | 1
| 1 | 2 | < 1 | < 0.5 | < 2 | < 1 | < 0.5 | [14 empty] | 9
| 1 | 3 | < 1 | < 0.5 | < 2 | < 1 | < 0.5 | [14 empty] | 16
| 1 | 4 | < 1 | 0.5 | < 2 | < 1 | < 0.5 | [14 empty] | < 1
| 1 | 5 | < 1 | < 0.5 | < 2 | < 1 | < 0.5 | [14 empty] | 15
| 1 | 6 | < 1 | 0.5 | < 2 | < 1 | < 0.5
| Mean |  |  | 0.5
| Median |  |  | 0.5
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
PDM3
| Z-Score (Absolute) |  | 0 | 0.67 | 2.02 | 0 | 1.01
| Analytical results for Th in OREAS 45f (Certified Value 9.99 ppm) | [20 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 10 | 13 | 14 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 10.3 | 8.624 | 10.7 | 9.8 | 9 | 9.7 | 10.9 | 8.6 | 10.3865 | 11.2 | 10.5 | 10.9 | 10.227 | [6 empty] | 1
| 1 | 2 | 10.6 | 8.661 | 10.4 | 9 | 9 | 9.6 | 10.7 | 8.6 | 10.239 | 10.9 | 10.3 | 10.5 | 9.68 | [6 empty] | #N/A
| 1 | 3 | 10.2 | 8.3278 | 10.4 | 9.2 | 10 | 9.9 | 11.1 | 9.3 | 9.8716 | 11.9 | 9.7 | 10.8 | 10.151 | [6 empty] | 16
| 1 | 4 | 10.6 | 8.4568 | 10.6 | 9.1 | 9.5 | 10.1 | 11 | 9.4 | 9.6433 | 11 | 9.9 | 10.7 | 9.06 | [6 empty] | 9.9905
| 1 | 5 | 10.3 | 7.227 | 10.6 | 9 | 9.5 | 9.8 | 11.3 | 9 | 10.1305 | 11 | 10.2 | 11.3 | 10.044 | [6 empty] | 73
| 1 | 6 | 10.3 | 8.3 | 10.7 | 9.7 | 9 | 9.8 | 11.2 | 9.3 | 10.1508 | 10.9 | 10.9 | 10.6 | 9.73
| Mean |  | 10.3833 | 8.2661 | 10.5667 | 9.3 | 9.3333 | 9.8167 | 11.0333 | 9.0333 | 10.0703 | 11.15 | 10.25 | 10.8 | 9.8153
| Median |  | 10.3 | 8.3923 | 10.6 | 9.15 | 9.25 | 9.8 | 11.05 | 9.15 | 10.1406 | 11 | 10.25 | 10.75 | 9.887
| Std Dev. |  | 0.17224 | 0.53011 | 0.13663 | 0.35777 | 0.40825 | 0.17224 | 0.21602 | 0.36148 | 0.26847 | 0.38341 | 0.42778 | 0.28284 | 0.43141
| Rel.Std.Dev. |  | 1.659% | 6.413% | 1.293% | 3.847% | 4.374% | 1.755% | 1.958% | 4.002% | 2.666% | 3.439% | 4.174% | 2.619% | 4.395%
| PDM3 |  | 3.932% | -17.26% | 5.767% | -6.911% | -6.578% | -1.74% | 10.44% | -9.581% | 0.7987% | 11.61% | 2.598% | 8.103% | -1.753%
| Z-Score (Absolute) |  | 0.29 | 1.67 | 0.46 | 0.71 | 0.68 | 0.23 | 0.89 | 0.96 | 0 | 1 | 0.17 | 0.67 | 0.24
| Analytical results for Ti in OREAS 45f (Certified Value 1.14 wt.%) | [20 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 36 | 39 | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES | BF*OES/MS | PF*OES | BF*OES | BF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES | 2
3
| 1 | 1 | 1.1112 | 1.097 | 1.18 | 1.1765 | 1.19 | 1.21 | 1.155 | 1.118 |  | 1.18 | 1.163 | 1.163 | 1.19 | 1.1707 | 1.14 | 1.15 | 1.157 | 1.0611 | 1.1219 | 1
| 1 | 2 | 1.1054 | 1.097 | 1.17 | 1.1735 | 1.21 | 1.23 | 1.105 | 1.1387 |  | 1.18 | 1.157 | 1.181 | 1.15 | 1.1478 | 1.14 | 1.15 | 1.121 | 1.0611 | 1.1218 | #N/A
| 1 | 3 | 1.0912 | 1.085 | 1.17 | 1.1687 | 1.2 | 1.19 | 1.14 | 1.1659 |  | 1.18 | 1.163 | 1.145 | 1.12 | 1.0366 | 1.11 | 1.14 | 1.151 | 1.127 | 1.0788 | 16
| 1 | 4 | 1.0932 | 1.097 | 1.17 | 1.1738 | 1.21 | 1.2 | 1.155 | 1.1283 |  | 1.17 | 1.163 | 1.163 | 1.18 | 1.0895 | 1.13 | 1.15 | 1.121 | 1.127 | 1.1123 | 1.1439
| 1 | 5 | 1.1025 | 1.127 | 1.18 | 1.2008 | 1.18 | 1.25 | 1.125 | 1.1452 |  | 1.17 | 1.163 | 1.175 | 1.16 | 1.0363 | 1.13 | 1.13 | 1.169 | 1.0731 | 1.1067 | 74
| 1 | 6 | 1.0928 | 1.097 | 1.15 | 1.1732 | 1.2 | 1.2 | 1.145 | 1.1383 |  | 1.18 | 1.163 | 1.157 | 1.15 | 1.045 | 1.12 | 1.16 | 1.115 | 1.0731 | 1.0857
| Mean |  | 1.0994 | 1.1 | 1.17 | 1.1777 | 1.1983 | 1.2133 | 1.1375 | 1.1391 |  | 1.1767 | 1.162 | 1.164 | 1.1583 | 1.0877 | 1.1283 | 1.1467 | 1.139 | 1.0871 | 1.1045
| Median |  | 1.0979 | 1.097 | 1.17 | 1.1736 | 1.2 | 1.205 | 1.1425 | 1.1385 |  | 1.18 | 1.163 | 1.163 | 1.155 | 1.0673 | 1.13 | 1.15 | 1.136 | 1.0731 | 1.1095
| Std Dev. |  | 0.0081808 | 0.014071 | 0.010954 | 0.01156 | 0.01169 | 0.022509 | 0.019429 | 0.016236 |  | 0.005164 | 0.0024495 | 0.012822 | 0.024833 | 0.059296 | 0.01169 | 0.010328 | 0.022768 | 0.031394 | 0.018318
| Rel.Std.Dev. |  | 0.7441% | 1.279% | 0.9363% | 0.9816% | 0.9756% | 1.855% | 1.708% | 1.425% |  | 0.4389% | 0.2108% | 1.102% | 2.144% | 5.452% | 1.036% | 0.9007% | 1.999% | 2.888% | 1.658%
| PDM3 |  | -3.888% | -3.835% | 2.284% | 2.96% | 4.761% | 6.072% | -0.5571% | -0.4202% |  | 2.867% | 1.585% | 1.76% | 1.264% | -4.913% | -1.359% | 0.2442% | -0.426% | -4.966% | -3.44%
| Z-Score (Absolute) |  | 0.96 | 0.95 | 0.6 | 0.77 | 1.23 | 1.56 | 0.12 | 0.08 | NA | 0.75 | 0.42 | 0.47 | 0.34 | 1.22 | 0.32 | 0.08 | 0.09 | 1.24 | 0.85
| Analytical results for Tl in OREAS 45f (Indicative Value < 0.5 ppm) | [20 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 05 | 13 | 14 | 19 | 24 | 25 | 33 | 34 | 39 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [10 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | 0.21 | < 2 | < 0.5 | < 0.5 | < 0.25 | < 0.5 | 0.177 | [10 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | 0.17 | < 2 | < 0.5 | < 0.5 | < 0.25 | < 0.5 | 0.152 | [10 empty] | 10
| 1 | 3 | < 0.5 | < 0.5 | 0.18 | < 2 | < 0.5 | < 0.5 | < 0.25 | < 0.5 | 0.147 | [10 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | 0.17 | < 2 | < 0.5 | < 0.5 | < 0.25 | < 0.5 | 0.141 | [10 empty] | < 0.5
| 1 | 5 | < 0.5 | < 0.5 | 0.13 | < 2 | < 0.5 | < 0.5 | < 0.25 | < 0.5 | 0.163 | [10 empty] | 16
| 1 | 6 | < 0.5 | < 0.5 | 0.22 | < 2 | < 0.5 | < 0.5 | < 0.25 | < 0.5 | 0.15
| Mean |  |  |  | 0.18 | [5 empty] | 0.155
| Median |  |  |  | 0.175 | [5 empty] | 0.151
| Std Dev. |  |  |  | 0.032249 | [5 empty] | 0.012977
| Rel.Std.Dev. |  |  |  | 17.92% | [5 empty] | 8.372%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Tm in OREAS 45f (Certified Value 0.26 ppm) | [20 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.26 | 0.22538 | 0.24 | 0.2 | 0.26 | 0.32 | 0.295 | < 0.25 | 0.25 | 0.231 | [9 empty] | 1
| 1 | 2 | 0.26 | 0.23192 | 0.27 | 0.2 | 0.27 | 0.33 | 0.298 | < 0.25 | 0.29 | 0.234 | [9 empty] | #N/A
| 1 | 3 | 0.24 | 0.21225 | 0.22 | 0.2 | 0.27 | 0.32 | 0.271 | < 0.25 | 0.29 | 0.216 | [9 empty] | 16
| 1 | 4 | 0.27 | 0.21504 | 0.21 | 0.2 | 0.3 | 0.33 | 0.271 | < 0.25 | 0.28 | 0.214 | [9 empty] | 0.26176
| 1 | 5 | 0.24 | 0.20858 | 0.24 | 0.2 | 0.29 | 0.33 | 0.272 | < 0.25 | 0.27 | 0.249 | [9 empty] | 75
| 1 | 6 | 0.26 | 0.20616 | 0.21 | 0.2 | 0.28 | 0.33 | 0.279 | < 0.25 | 0.29 | 0.215
| Mean |  | 0.255 | 0.21656 | 0.23167 | 0.2 | 0.27833 | 0.32667 | 0.281 |  | 0.27833 | 0.2265
| Median |  | 0.26 | 0.21365 | 0.23 | 0.2 | 0.275 | 0.33 | 0.2755 |  | 0.285 | 0.2235
| Std Dev. |  | 0.012247 | 0.010064 | 0.023166 | 3.0405e-17 | 0.01472 | 0.005164 | 0.01241 |  | 0.016021 | 0.014011
| Rel.Std.Dev. |  | 4.803% | 4.647% | 10% | 1.52e-14% | 5.288% | 1.581% | 4.416% |  | 5.756% | 6.186%
| PDM3 |  | -2.581% | -17.27% | -11.5% | -23.59% | 6.333% | 24.8% | 7.351% |  | 6.333% | -13.47%
| Z-Score (Absolute) |  | 0 | 1 | 0.61 | NA | 0.61 | 1.86 | 0.67 | 3.37 | 0.61 | 0.74
Results from laboratory 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 45f (Certified Value 2.32 ppm) | [20 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 10 | 13 | 14 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [6 empty] | 2
3
| 1 | 1 | 2.25 | 2.2343 | 2.6 | 2.2 | 2.5 | 2.3 | 2.5 | 2.2 | 2.4049 | 2.5 | 2.4 | 2.47 | 2.032 | [6 empty] | 1
| 1 | 2 | 2.33 | 2.2939 | 2.4 | 2.2 | 2.5 | 2.1 | 2.4 | 2.2 | 2.4571 | 2.5 | 2.3 | 2.3 | 1.889 | [6 empty] | #N/A
| 1 | 3 | 2.24 | 2.1723 | 2.4 | 2 | 3 | 2.1 | 2.6 | 2.2 | 2.187 | 2.6 | 2.3 | 2.4 | 2.02 | [6 empty] | 16
| 1 | 4 | 2.32 | 2.2352 | 2.6 | 2.2 | 2.5 | 2.2 | 2.7 | 2.2 | 2.2931 | 2.45 | 2.2 | 2.31 | 1.821 | [6 empty] | 2.3179
| 1 | 5 | 2.26 | 2.1434 | 2.5 | 2.2 | 2.5 | 2.1 | 2.6 | 2.3 | 2.2497 | 2.45 | 2.4 | 2.62 | 2.034 | [6 empty] | 76
| 1 | 6 | 2.25 | 2.1682 | 2.5 | 2.2 | 3 | 2.3 | 2.5 | 2.2 | 2.2476 | 2.53 | 2.6 | 2.44 | 1.797
| Mean |  | 2.275 | 2.2079 | 2.5 | 2.1667 | 2.6667 | 2.1833 | 2.55 | 2.2167 | 2.3066 | 2.505 | 2.3667 | 2.4233 | 1.9322
| Median |  | 2.255 | 2.2033 | 2.5 | 2.2 | 2.5 | 2.15 | 2.55 | 2.2 | 2.2714 | 2.5 | 2.35 | 2.42 | 1.9545
| Std Dev. |  | 0.03937 | 0.056295 | 0.089443 | 0.08165 | 0.2582 | 0.098319 | 0.10488 | 0.040825 | 0.10344 | 0.056125 | 0.13663 | 0.1181 | 0.11004
| Rel.Std.Dev. |  | 1.731% | 2.55% | 3.578% | 3.768% | 9.682% | 4.503% | 4.113% | 1.842% | 4.485% | 2.241% | 5.773% | 4.873% | 5.695%
| PDM3 |  | -1.853% | -4.748% | 7.854% | -6.526% | 15.04% | -5.807% | 10.01% | -4.369% | -0.4913% | 8.07% | 2.102% | 4.547% | -16.64%
| Z-Score (Absolute) |  | 0.17 | 0.54 | 1.06 | 0.77 | 1.97 | 0.67 | 1.33 | 0.49 | 0 | 1.09 | 0.33 | 0.64 | 2.05
| Analytical results for V in OREAS 45f (Certified Value 261 ppm) | [20 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 05 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 |  |  |  |  | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES | BF*OES | BF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES |  |  |  |  | 0
0
| 1 | 1 | 259 | 260 | 250 | 260 | 307 |  | 250 | 262 | 232 | 277 | 270.9973 | 258 | 270 | 271 | 241.5888 |  |  |  |  | 1
| 1 | 2 | 258 | 265 | 250 | 270 | 302 |  | 250 | 264 | 232 | 330 | 268.2836 | 262 | 275 | 267 | 240.5196 |  |  |  |  | #N/A
| 1 | 3 | 261 | 265 | 300 | 270 | 304 |  | 250 | 260 | 227 | 298 | 255.221 | 263 | 265 | 272 | 237.1812 |  |  |  |  | 16
| 1 | 4 | 260 | 265 | 250 | 260 | 309 |  | 250 | 264 | 233 | 271 | 259.0043 | 260 | 280 | 268 | 242.904 |  |  |  |  | 261.4327
| 1 | 5 | 260 | 265 | 300 | 270 | 295 |  | 250 | 261 | 235 | 262 | 269.1703 | 264 | 270 | 251 | 240.6492 |  |  |  |  | 77
| 1 | 6 | 264 | 260 | 300 | 270 | 314 |  | 250 | 262 | 230 | 295 | 263.2829 | 262 | 270 | 257 | 230.634
| Mean |  | 260.3333 | 263.3333 | 275 | 266.6667 | 305.1667 |  | 250 | 262.1667 | 231.5 | 288.8333 | 264.3266 | 261.5 | 271.6667 | 264.3333 | 238.9128
| Median |  | 260 | 265 | 275 | 270 | 305.5 |  | 250 | 262 | 232 | 286 | 265.7833 | 262 | 270 | 267.5 | 240.5844
| Std Dev. |  | 2.0656 | 2.582 | 27.3861 | 5.164 | 6.4936 |  | 0 | 1.6021 | 2.7386 | 24.4738 | 6.2601 | 2.1679 | 5.164 | 8.4301 | 4.4772
| Rel.Std.Dev. |  | 0.7934% | 0.9805% | 9.959% | 1.936% | 2.128% |  | 0% | 0.6111% | 1.183% | 8.473% | 2.368% | 0.829% | 1.901% | 3.189% | 1.874%
| PDM3 |  | -0.4205% | 0.727% | 5.19% | 2.002% | 16.73% |  | -4.373% | 0.2808% | -11.45% | 10.48% | 1.107% | 0.02576% | 3.915% | 1.11% | -8.614%
| Z-Score (Absolute) |  | 0.42 | 0.06 | 1.33 | 0.34 | 4.92 | NA | 1.65 | 0.2 | 3.85 | 2.98 | 0.06 | 0.28 | 0.93 | 0.06 | 2.97
| Analytical results for W in OREAS 45f (Indicative Value 1.37 ppm) | [20 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 05 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [6 empty] | 2
2
| 1 | 1 | 1 | 1 | 1.6 |  | < 5 | 1.6 | < 0.5 | 4 | 1.2296 | 1.5 | 2 | 1.2 | 1.19 | [6 empty] | 1
| 1 | 2 | 1 | 2 | 1.6 |  | < 5 | 1.6 | < 0.5 | 2 | < 1 | 1.54 | < 2 | 4.5 | 1.07 | [6 empty] | 11
| 1 | 3 | 1 | 1 | 1.6 |  | < 5 | 1.3 | < 0.5 | 4 | 1.0491 | 1.54 | < 2 | 2 | 1.17 | [6 empty] | 16
| 1 | 4 | 1 | 1 | 1.6 |  | < 5 | 1.9 | < 0.5 | 2 | 1.1301 | 1.49 | < 2 | 1.7 | 1.05 | [6 empty] | 1.374
| 1 | 5 | 2 | 2 | 1.2 |  | < 5 | 0.8 | 1 | 2 | < 1 | 1.53 | < 2 | 2 | 1.81 | [6 empty] | 17
| 1 | 6 | 1 | 2 | 1.4 |  | < 5 | 1.3 | < 0.5 | 4 | 1.1301 | 1.46 | < 2 | 1.9 | 1.19
| Mean |  | 1.1667 | 1.5 | 1.5 |  |  | 1.4167 | 1 | 3 | 1.1347 | 1.51 | 2 | 2.2167 | 1.2467
| Median |  | 1 | 1.5 | 1.6 |  |  | 1.45 | 1 | 3 | 1.1301 | 1.515 | 2 | 1.95 | 1.18
| Std Dev. |  | 0.40825 | 0.54772 | 0.16733 |  |  | 0.37639 |  | 1.0954 | 0.07387 | 0.032249 |  | 1.1583 | 0.28268
| Rel.Std.Dev. |  | 34.99% | 36.51% | 11.16% |  |  | 26.57% |  | 36.51% | 6.51% | 2.136% |  | 52.25% | 22.67%
| PDM3 |  | -15.09% | 9.173% | 9.173% |  |  | 3.108% | -27.22% | 118.3% | -17.41% | 9.901% | 45.56% | 61.33% | -9.265%
| Z-Score (Absolute) |  | NA | NA | 0.08 | NA | 1.88 | 0.08 | 1.96 | NA | 0.97 | 0.09 | NA | 1.37 | 0.38
Results from laboratories 5, 13, 23 & 33 were removed due to their1 ppm reading resolution.
| Analytical results for Y in OREAS 45f (Certified Value 14.7 ppm) | [20 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 04 | 05 | 10 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 33 | 34 | 39 |  |  |  |  | ppm
|  |  | BF*OES | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES/MS | PF*MS | BF*OES | BF*OES | PF*MS | PF*MS | PF*MS | PF*MS | BF*MS | PF*OES |  |  |  |  | 1
2
| 1 | 1 | 16 | 15.7 | 25.6357 | 15.1 | 15.7 |  | 13 | 13.9 | 14 | 13 | 13.7597 | 16.9 | 16.9 | 15.37 | 12.305 |  |  |  |  | 1
| 1 | 2 | 14 | 15.7 | 25.8625 | 15 | 14.8 |  | 13 | 14.4 | 14 | 14.1 | 13.8785 | 17.1 | 16.4 | 15.59 | 12.114 |  |  |  |  | #N/A
| 1 | 3 | 15 | 15.5 | 25.3205 | 14.8 | 16.6 |  | 14 | 13.8 | 13 | 13.5 | 13.8278 | 16 | 15.5 | 16.08 | 11.689 |  |  |  |  | 16
| 1 | 4 | 14 | 16.5 | 25.2208 | 15.9 | 16.3 |  | 13 | 15 | 14 | 14.4 | 13.6997 | 17.1 | 15.1 | 15.77 | 11.94 |  |  |  |  | 14.7068
| 1 | 5 | 16 | 16.4 | 23.6386 | 16.4 | 14.8 |  | 14 | 12.3 | 13 | 14.2 | 13.9489 | 16.5 | 16.1 | 14.83 | 11.852 |  |  |  |  | 78
| 1 | 6 | 15 | 15.8 | 25.2739 | 15.4 | 16.4 |  | 14 | 15.3 | 13 | 13.4 | 13.6883 | 16.9 | 16.6 | 15.36 | 11.325
| Mean |  | 15 | 15.9333 | 25.1587 | 15.4333 | 15.7667 |  | 13.5 | 14.1167 | 13.5 | 13.7667 | 13.8005 | 16.75 | 16.1 | 15.5 | 11.8708
| Median |  | 15 | 15.75 | 25.2972 | 15.25 | 16 |  | 13.5 | 14.15 | 13.5 | 13.8 | 13.7937 | 16.9 | 16.25 | 15.48 | 11.896
| Std Dev. |  | 0.89443 | 0.41312 | 0.78444 | 0.60882 | 0.80664 |  | 0.54772 | 1.0685 | 0.54772 | 0.5465 | 0.10322 | 0.42778 | 0.68411 | 0.42492 | 0.34169
| Rel.Std.Dev. |  | 5.963% | 2.593% | 3.118% | 3.945% | 5.116% |  | 4.057% | 7.569% | 4.057% | 3.97% | 0.7479% | 2.554% | 4.249% | 2.741% | 2.878%
| PDM3 |  | 1.994% | 8.34% | 71.07% | 4.94% | 7.207% |  | -8.206% | -4.012% | -8.206% | -6.392% | -6.162% | 13.89% | 9.473% | 5.394% | -19.28%
| Z-Score (Absolute) |  | 0.12 | 0.38 | 5.33 | 0.12 | 0.29 | NA | 0.92 | 0.59 | 0.92 | 0.78 | 0.76 | 0.82 | 0.47 | 0.15 | 1.79
| Analytical results for Yb in OREAS 45f (Certified Value 1.79 ppm) | [20 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 05 | 10 | 14 | 19 | 20 | 21 | 25 | 33 | 34 | 39 | [9 empty] | ppm
|  |  | PF*OES/MS | BF*OES/MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | [9 empty] | 2
3
| 1 | 1 | 1.8 | 1.0924 | 1.81 | 2 | 1.96 | 1.82 | 1.8 | 1.3 | 1.9 | 1.638 | [9 empty] | 1
| 1 | 2 | 1.7 | 1.1364 | 2.1 | 2 | 1.88 | 1.88 | 1.9 | 1.4 | 1.8 | 1.619 | [9 empty] | #N/A
| 1 | 3 | 1.6 | 1.0944 | 1.58 | 1.5 | 1.93 | 1.9 | 1.73 | 1.3 | 2 | 1.512 | [9 empty] | 16
| 1 | 4 | 1.8 | 1.0984 | 1.89 | 1.5 | 1.88 | 1.9 | 1.94 | 1.2 | 1.9 | 1.505 | [9 empty] | 1.7942
| 1 | 5 | 1.6 | 1.0635 | 1.69 | 2 | 1.98 | 1.89 | 1.78 | 1.3 | 1.7 | 1.634 | [9 empty] | 79
| 1 | 6 | 1.7 | 1.05 | 1.72 | 2 | 1.87 | 1.87 | 1.88 | 1.3 | 1.7 | 1.367
| Mean |  | 1.7 | 1.0892 | 1.7983 | 1.8333 | 1.9167 | 1.8767 | 1.8383 | 1.3 | 1.8333 | 1.5458
| Median |  | 1.7 | 1.0934 | 1.765 | 2 | 1.905 | 1.885 | 1.84 | 1.3 | 1.85 | 1.5655
| Std Dev. |  | 0.089443 | 0.030163 | 0.1817 | 0.2582 | 0.046762 | 0.030111 | 0.080602 | 0.063246 | 0.12111 | 0.10622
| Rel.Std.Dev. |  | 5.261% | 2.769% | 10.1% | 14.08% | 2.44% | 1.604% | 4.385% | 4.865% | 6.606% | 6.871%
| PDM3 |  | -5.252% | -39.3% | 0.2287% | 2.179% | 6.824% | 4.595% | 2.458% | -27.55% | 2.179% | -13.84%
| Z-Score (Absolute) |  | 0.97 | 6.06 | 0.15 | 0.15 | 0.84 | 0.51 | 0.19 | 4.3 | 0.15 | 2.25
| Analytical results for Zn in OREAS 45f (Certified Value 34.3 ppm) | [20 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 04 | 05 | 10 | 12 | 13 | 14 | 17 | 19 | 24 | 25 | 33 | 34 | 39 | [6 empty] | ppm
|  |  | BF*OES | PF*OES/MS | BF*OES/MS | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES | [6 empty] | 1
1
| 1 | 1 | 37 | 34 | 30 | 50 | < 100 | 40 |  | < 50 | 36.4327 | 39.2 | 40 | 20 | 28.523 | [6 empty] | 1
| 1 | 2 | 42 | 35 | 25 | 50 | < 100 | 30 |  | < 50 | 32.4599 | 36.8 | 40 | 31 | 29 | [6 empty] | 1
| 1 | 3 | 39 | 34 | 35 | < 50 | < 100 | 40 |  | 50 | 35.7126 | 37.1 | 40 | 43 | 28.139 | [6 empty] | 16
| 1 | 4 | 38 | 34 | 25 | 50 | < 100 | 40 |  | < 50 | 34.3864 | 35.3 | 40 | 35 | 27.62 | [6 empty] | 34.3387
| 1 | 5 | 40 | 35 | 25 | < 50 | < 100 | 30 |  | < 50 | 33.3379 | 40.4 | 40 | 59 | 32.411 | [6 empty] | 80
| 1 | 6 | 39 | 33 | 30 | < 50 | < 100 | 30 |  | < 50 | 34.3904 | 37.3 | 35 | < 5 | 30
| Mean |  | 39.1667 | 34.1667 | 28.3333 | 50 |  | 35 |  | 50 | 34.4533 | 37.6833 | 39.1667 | 37.6 | 29.2822
| Median |  | 39 | 34 | 27.5 | 50 |  | 35 |  | 50 | 34.3884 | 37.2 | 40 | 35 | 28.7615
| Std Dev. |  | 1.7224 | 0.75277 | 4.0825 | 0 |  | 5.4772 |  |  | 1.4649 | 1.8236 | 2.0412 | 14.5534 | 1.7338
| Rel.Std.Dev. |  | 4.398% | 2.203% | 14.41% | 0% |  | 15.65% |  |  | 4.252% | 4.839% | 5.212% | 38.71% | 5.921%
| PDM3 |  | 14.06% | -0.501% | -17.49% | 45.61% |  | 1.926% |  | 45.61% | 0.3338% | 9.74% | 14.06% | 9.497% | -14.73%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 0.88 | NA | 2.22 | NA | NA | NA | 0 | 0.46 | 0.67 | 0.39 | 0.74
Results from laboratories 12, 14 & 19 were removed due to their 10 ppm or greater reading resolution.
| Analytical results for Zr in OREAS 45f (Certified Value 261 ppm) | [20 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 04 | 05 | 10 | 17 | 20 | 21 | 25 | 34 | [11 empty] | ppm
|  |  | BF*OES | PF*OES/MS | BF*OES/MS | BF*OES/MS | BF*OES | BF*MS | PF*MS | BF*OES | [11 empty] | 0
0
| 1 | 1 | 233 | 252 | 265 |  | 244.8 | 290.6 | 266 | 265 | [11 empty] | 1
| 1 | 2 | 223 | 268 | 260 |  | 247.6 | 301.4 | 260 | 255 | [11 empty] | #N/A
| 1 | 3 | 236 | 266 | 260 |  | 251 | 292.7 | 265 | 267 | [11 empty] | 16
| 1 | 4 | 224 | 299 | 265 |  | 252.2 | 295.8 | 269 | 264 | [11 empty] | 261.2643
| 1 | 5 | 236 | 280 | 270 |  | 239.1 | 290 | 270 | 237 | [11 empty] | 81
| 1 | 6 | 230 | 275 | 270 |  | 239.2 | 277.7 | 274 | 247
| Mean |  | 230.3333 | 273.3333 | 265 |  | 245.65 | 291.3667 | 267.3333 | 255.8333
| Median |  | 231.5 | 271.5 | 265 |  | 246.2 | 291.65 | 267.5 | 259.5
| Std Dev. |  | 5.7504 | 15.7692 | 4.4721 |  | 5.6667 | 7.8909 | 4.8028 | 11.9066
| Rel.Std.Dev. |  | 2.497% | 5.769% | 1.688% |  | 2.307% | 2.708% | 1.797% | 4.654%
| PDM3 |  | -11.84% | 4.619% | 1.43% |  | -5.976% | 11.52% | 2.323% | -2.079%
| Z-Score (Absolute) |  | 2.55 | 0.61 | 0 | NA | 1.42 | 1.94 | 0.17 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C551: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C570: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C608: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C627: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C645: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C663: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C681: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C717: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C736: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C754: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C772: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C790: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C808: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C826: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C844: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C862: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C880: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1045: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1154: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1173: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1191: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1228: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 45f (Indicative Value 0.088 ppm) | [32 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | Various [a] | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES |  |  | 3
|  |  | 21: see footer |  |  | 3
| 1 | 1 | < 0.25 | < 0.2 | 0.11 | 0.2 | 18.4901 | < 0.5 | 0.06 | 0.079 | 0.5 | 0.7 | 0.2 | < 0.2 | < 0.1 | 0.5 | 0.15 | 0.095424 | 0.088 | 0.4 | 0.08 | 0.062 | 0.08 | 0.078 | < 0.5 | 0.4 | 0.073 | 0.47 | 0.5 | 0.093333 | 0.968 |  |  | 1
| 1 | 2 | < 0.25 | < 0.2 | 0.15 | 0.1 | 18.2856 | < 0.5 | 0.03 | 0.067 | 0.5 | 0.83 | 0.2 | < 0.2 | < 0.1 | 0.8 | 0.15 | 0.083465 | 0.095 | 0.4 | 0.069 | 0.084 | 0.08 | 0.076 | < 0.5 | 0.17 | 0.08 | 0.42 | 0.6 | 0.063333 | 1.009 |  |  | 7
| 1 | 3 | < 0.25 | < 0.2 | 0.13 | 0.1 | 18.5086 | < 0.5 | 0.04 | 0.069 | 0.6 | 0.67 | 0.2 | < 0.2 | < 0.1 | 0.3 | 0.17 | 0.085954 | 0.105 | 0.4 | 0.077 | 0.073 | 0.07 | 0.073 | < 0.5 | 0.21 | 0.083 | 0.39 | < 0.5 | 0.13333 | 1.05 |  |  | 16
| 1 | 4 | < 0.25 | < 0.2 | 0.15 | 0.1 | 17.8675 | < 0.5 | 0.06 | 0.076 | 0.4 | 0.63 | 0.2 | < 0.2 | < 0.1 | 0.6 | 0.14 | 0.092634 | 0.073 | 0.4 | 0.074 | 0.074 | 0.08 | 0.08 | < 0.5 | 0.35 | 0.073 | 0.4 | < 0.5 | 0.12333 | 0.938 |  |  | 0.088056
| 1 | 5 | < 0.25 | < 0.2 | 0.15 | 0.1 | 19.6315 | < 0.5 | 0.03 | 0.076 | 0.4 | 0.76 | 0.2 | < 0.2 | < 0.1 | 0.6 | 0.14 | 0.085054 | 0.101 | 0.4 | 0.068 | 0.072 | 0.05 | 0.071 | < 0.5 | 0.4 | 0.073 | 0.4 | < 0.5 | 0.096667 | 0.96 |  |  | 21
| 1 | 6 | < 0.25 | < 0.2 | 0.11 | 0.1 | 18.1905 | < 0.5 | 0.06 | 0.075 | 0.5 | 0.53 | 0.2 | < 0.2 | < 0.1 | 0.4 | 0.14 | 0.092837 | 0.087 | 0.3 | 0.083 | 0.073 | 0.1 | 0.073 | < 0.5 | 0.37 | 0.085 | 0.39 | < 0.5 | 0.066667 | 0.989
| Mean |  |  |  | 0.13333 | 0.11667 | 18.4956 |  | 0.046667 | 0.073667 | 0.48333 | 0.68667 | 0.2 |  |  | 0.53333 | 0.14833 | 0.089228 | 0.0915 | 0.38333 | 0.075167 | 0.073 | 0.076667 | 0.075167 |  | 0.31667 | 0.077833 | 0.41167 | 0.55 | 0.096111 | 0.98567
| Median |  |  |  | 0.14 | 0.1 | 18.3879 |  | 0.05 | 0.0755 | 0.5 | 0.685 | 0.2 |  |  | 0.55 | 0.145 | 0.089294 | 0.0915 | 0.4 | 0.0755 | 0.073 | 0.08 | 0.0745 |  | 0.36 | 0.0765 | 0.4 | 0.55 | 0.095 | 0.9785
| Std Dev. |  |  |  | 0.019664 | 0.040825 | 0.60369 |  | 0.015055 | 0.0046332 | 0.075277 | 0.10405 | 3.0405e-17 |  |  | 0.17512 | 0.01169 | 0.0049875 | 0.011485 | 0.040825 | 0.0059805 | 0.0069857 | 0.01633 | 0.0034303 |  | 0.10073 | 0.0055287 | 0.030605 | 0.070711 | 0.028551 | 0.039843
| Rel.Std.Dev. |  |  |  | 14.75% | 34.99% | 3.264% |  | 32.26% | 6.289% | 15.57% | 15.15% | 1.52e-14% |  |  | 32.83% | 7.881% | 5.59% | 12.55% | 10.65% | 7.956% | 9.569% | 21.3% | 4.564% |  | 31.81% | 7.103% | 7.434% | 12.86% | 29.71% | 4.042%
| PDM3 |  |  |  | 51.42% | 32.49% | 20904% |  | -47% | -16.34% | 448.9% | 679.8% | 127.1% |  |  | 505.7% | 68.45% | 1.331% | 3.911% | 335.3% | -14.64% | -17.1% | -12.93% | -14.64% |  | 259.6% | -11.61% | 367.5% | 524.6% | 9.148% | 1019%
| Z-Score (Absolute) |  | 0.62 | 0 | 0.83 | NA | NA | 3.75 | 1.33 | 0.66 | NA | 14.65 | NA | 0 | 1.25 | NA | 1.21 | 0.27 | 0.21 | NA | 0.62 | 0.67 | 0.58 | 0.62 | 3.75 | 5.41 | 0.55 | 7.78 | 6.24 | 0.1 | 22.12
Laboratory10's results were excluded due to extreme high bias. Results from laboratories 7, 15, 18, 21 & 27 were removed due to their 0.1 ppm reading resolution.
[a]: Lab 21: Rnd 01 - 4A*MS, Rnd 01 - 4A*OES
| Analytical results for Al in OREAS 45f (Certified Value 10.16 wt.%) | [32 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  |  | 2
3
| 1 | 1 | 10.4905 | 10.11 | 10.5 | 10.7 | 10.8 | > 10 | 9.94 | 8.69 | 8.77 | 9.89 | 10.2 | 9.96 | 10.7047 | 10.2076 | 9.95 | 11.5 | 9.65 | 10.05 | 10.65 | 10.6 | 9.8 | 9.78 | 10.65 | 10.5 | 9.94 | 9.7179 | 10.3519 | 9.784 |  |  |  | 1
| 1 | 2 | 10.451 | 10.3 | 10.4 | 10.7 | 10.4 | > 10 | 9.82 | 5.09 | 9.3 | 9.77 | 9.62 | 8.97 | 10.6059 | 10.2181 | 9.88 | 11 | 9.89 | 10 | 11 | 10.1 | 10.3 | 10 | 11.05 | 10.4 | 9.78 | 9.5565 | 10.4428 | 9.9433 |  |  |  | #N/A
| 1 | 3 | 10.3975 | 9.81 | 10.4 | 10.6 | 10.3 | > 10 | 9.68 | 7.3 | 8.76 | 10 | 10.13 | 9.55 | 10.5792 | 10.124 | 9.91 | 11.3 | 9.87 | 10.3 | 10.45 | 10.55 | 9.51 | 9.23 | 10.75 | 10.4 | 9.82 | 9.5647 | 10.257 | 10.0244 |  |  |  | 16
| 1 | 4 | 10.5432 | 10.55 | 10.5 | 10.4 | 10.4 | > 10 | 9.82 | 5.82 | 8.77 | 9.2 | 10.14 | 9.11 | 10.7303 | 10.1341 | 9.86 | 11.5 | 9.61 | 9.9 | 10.95 | 10.2 | 9 | 9.29 | 10.65 | 10.2 | 9.73 | 9.5012 | 10.224 | 9.9136 |  |  |  | 10.1599
| 1 | 5 | 10.4819 | 10.22 | 11 | 10.4 | 10.6 | > 10 | 9.82 | 7.42 | 8.27 | 10.6 | 10.26 | 9.63 | 10.7989 | 10.3146 | 9.96 | 10.7 | 9.76 | 10.15 | 10.45 | 9.72 | 10.3 | 9.84 | 10.65 | 10.2 | 9.95 | 9.7985 | 10.158 | 9.8488 |  |  |  | 83
| 1 | 6 | 10.4426 | 10.11 | 10.5 | 10.6 | 10.2 | 9.99 | 9.63 | 6.08 | 8.4 | 9.9 | 10.1 | 9.82 | 10.4502 | 10.2522 | 9.83 | 10.5 | 9.81 | 10.1 | 11.1 | 10.05 | 9.56 | 9.72 | 10.55 | 10.3 | 9.92 | 9.6113 | 10.2434 | 9.8978
| Mean |  | 10.4678 | 10.1833 | 10.55 | 10.5667 | 10.45 | 9.99 | 9.785 | 6.7333 | 8.7117 | 9.8933 | 10.075 | 9.5067 | 10.6449 | 10.2084 | 9.8983 | 11.0833 | 9.765 | 10.0833 | 10.7667 | 10.2033 | 9.745 | 9.6433 | 10.7167 | 10.3333 | 9.8567 | 9.625 | 10.2795 | 9.902
| Median |  | 10.4665 | 10.165 | 10.5 | 10.6 | 10.4 | 9.99 | 9.82 | 6.69 | 8.765 | 9.895 | 10.135 | 9.59 | 10.6553 | 10.2129 | 9.895 | 11.15 | 9.785 | 10.075 | 10.8 | 10.15 | 9.68 | 9.75 | 10.65 | 10.35 | 9.87 | 9.588 | 10.2502 | 9.9057
| Std Dev. |  | 0.049515 | 0.24476 | 0.22583 | 0.13663 | 0.21679 |  | 0.11203 | 1.3102 | 0.35986 | 0.44872 | 0.23011 | 0.39144 | 0.12516 | 0.072004 | 0.051153 | 0.42151 | 0.11485 | 0.13663 | 0.28752 | 0.33025 | 0.5027 | 0.31181 | 0.17512 | 0.12111 | 0.092664 | 0.11186 | 0.10156 | 0.081955
| Rel.Std.Dev. |  | 0.473% | 2.404% | 2.141% | 1.293% | 2.075% |  | 1.145% | 19.46% | 4.131% | 4.536% | 2.284% | 4.118% | 1.176% | 0.7053% | 0.5168% | 3.803% | 1.176% | 1.355% | 2.67% | 3.237% | 5.159% | 3.233% | 1.634% | 1.172% | 0.9401% | 1.162% | 0.9879% | 0.8277%
| PDM3 |  | 3.031% | 0.2308% | 3.84% | 4.004% | 2.856% | -1.672% | -3.69% | -33.73% | -14.25% | -2.624% | -0.8355% | -6.429% | 4.774% | 0.4779% | -2.574% | 9.089% | -3.887% | -0.7535% | 5.972% | 0.4277% | -4.084% | -5.084% | 5.48% | 1.707% | -2.984% | -5.264% | 1.178% | -2.538%
| Z-Score (Absolute) |  | 0.81 | 0.22 | 0.98 | 1.01 | 0.77 | 0.19 | 0.61 | 6.96 | 2.84 | 0.39 | 0.01 | 1.19 | 1.18 | 0.27 | 0.38 | 2.09 | 0.65 | 0.01 | 1.43 | 0.26 | 0.7 | 0.91 | 1.33 | 0.53 | 0.46 | 0.94 | 0.42 | 0.37
| Analytical results for As in OREAS 45f (Certified Value 9.67 ppm) | [32 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES |  |  | 2
3
| 1 | 1 | 9 | 12 | 10 | 11.9729 | 10 | 10.2 | 9.7 | 9.9 | 8 | 10 | 10 | 9 | 9 | 9.4 | 9.7076 | 9.38 | 8.849 | 10 | 9.5 | 9.4 | 9.7 | 9.76 | 9 | 10 | 9.91 | 13.9 | 14 | 10.3667 | 12.001 |  |  | 1
| 1 | 2 | 11 | 11 | 11 | 10.5648 | 9 | 10.2 | 9.6 | 9.8 | 9 | 10 | 10 | 8 | 8 | 9.3 | 9.3151 | 9.04 | 12.6821 | 10 | 9.17 | 9.11 | 10.5 | 9.08 | 8 | 13 | 10.15 | 12.1 | 13 | 10.7 | 11.217 |  |  | 1
| 1 | 3 | 9 | 11 | 11 | 10.2422 | 9 | 9.8 | 9 | 9.7 | 10 | 10 | 9 | 8 | 11 | 9.7 | 9.6794 | 9.78 | 9.9844 | 10 | 9.14 | 8.59 | 9.5 | 9.35 | 9 | 9 | 9.52 | 13.7 | 13 | 10.2 | 21.259 |  |  | 16
| 1 | 4 | 10 | 10 | 10 | 10.4314 | 10 | 9.9 | 9.64 | 10 | 6 | 11 | 9 | 8 | 13 | 9.7 | 9.3554 | 10.1 | 12.3276 | 10 | 9.5 | 9.18 | 8.1 | 8.96 | 9 | 9 | 9.1 | 11.6 | 14 | 9.7333 | 12.451 |  |  | 9.6675
| 1 | 5 | 10 | 11 | 10 | 10.2373 | 9 | 9.8 | 9.39 | 9.9 | 9 | 9 | 9 | 9 | 8 | 9.2 | 9.2179 | 10.1 | 10.1453 | 10 | 9.15 | 8.97 | 9.3 | 8.99 | 10 | 14 | 9.61 | 9.9 | 14 | 10.9 | 10.601 |  |  | 84
| 1 | 6 | 8 | 11 | 11 | 11.0735 | 10 | 9.6 | 9.15 | 9.2 | 9 | 9 | 9 | 9 | 8 | 9.3 | 9.9161 | 9.71 | 11.9466 | 10 | 10.05 | 9.22 | 11.1 | 9.19 | 10 | 10 | 9.9 | 12.1 | 13 | 9.7 | 15.342
| Mean |  | 9.5 | 11 | 10.5 | 10.7537 | 9.5 | 9.9167 | 9.4133 | 9.75 | 8.5 | 9.8333 | 9.3333 | 8.5 | 9.5 | 9.4333 | 9.5319 | 9.685 | 10.9891 | 10 | 9.4183 | 9.0783 | 9.7 | 9.2217 | 9.1667 | 10.8333 | 9.6983 | 12.2167 | 13.5 | 10.2667 | 13.8118
| Median |  | 9.5 | 11 | 10.5 | 10.4981 | 9.5 | 9.85 | 9.495 | 9.85 | 9 | 10 | 9 | 8.5 | 8.5 | 9.35 | 9.5174 | 9.745 | 11.0459 | 10 | 9.335 | 9.145 | 9.6 | 9.135 | 9 | 10 | 9.755 | 12.1 | 13.5 | 10.2833 | 12.226
| Std Dev. |  | 1.0488 | 0.63246 | 0.54772 | 0.67179 | 0.54772 | 0.24014 | 0.28605 | 0.2881 | 1.3784 | 0.75277 | 0.5164 | 0.54772 | 2.0736 | 0.21602 | 0.27457 | 0.41549 | 1.5412 | 0 | 0.35312 | 0.27752 | 1.0354 | 0.29983 | 0.75277 | 2.137 | 0.3707 | 1.4703 | 0.54772 | 0.49171 | 3.9994
| Rel.Std.Dev. |  | 11.04% | 5.75% | 5.216% | 6.247% | 5.766% | 2.422% | 3.039% | 2.955% | 16.22% | 7.655% | 5.533% | 6.444% | 21.83% | 2.29% | 2.881% | 4.29% | 14.02% | 0% | 3.749% | 3.057% | 10.67% | 3.251% | 8.212% | 19.73% | 3.822% | 12.03% | 4.057% | 4.789% | 28.96%
| PDM3 |  | -1.733% | 13.78% | 8.611% | 11.23% | -1.733% | 2.577% | -2.629% | 0.853% | -12.08% | 1.715% | -3.457% | -12.08% | -1.733% | -2.423% | -1.403% | 0.1807% | 13.67% | 3.439% | -2.578% | -6.095% | 0.3358% | -4.612% | -5.181% | 12.06% | 0.3186% | 26.37% | 39.64% | 6.197% | 42.87%
| Z-Score (Absolute) |  | 0.37 | 2.4 | 1.48 | 1.95 | 0.37 | 0.4 | 0.53 | 0.1 | 2.21 | 0.25 | 0.67 | 2.21 | 0.37 | 0.49 | 0.31 | 0.02 | 2.38 | 0.56 | 0.52 | 1.15 | 0 | 0.88 | 0.98 | 2.1 | 0 | 4.65 | 7.02 | 1.05 | 7.6
| Analytical results for B in OREAS 45f (Indicative Value 9.47 ppm) | [32 empty] | Indicative
B | Round | Replicate | 04: Lab |  |  | 1
| No | No | 04: 10 |  |  | ppm
|  |  | 04: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 04: 9.149 |  |  | 1
| 1 | 2 | 04: 9.7399 |  |  | 9
| 1 | 3 | 04: 9.655 |  |  | 16
| 1 | 4 | 04: 9.4662 |  |  | 9.4742
| 1 | 5 | 04: 9.5874 |  |  | 22
| 1 | 6 | 04: 9.2475
| Mean |  | 04: 9.4742
| Median |  | 04: 9.5268
| Std Dev. |  | 04: 0.23382
| Rel.Std.Dev. |  | 04: 2.468%
| PDM3 |  | 04: -3.331e-14%
Z-Score (Absolute) | bar: 04
| Analytical results for Ba in OREAS 45f (Certified Value 206 ppm) | [32 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 210.64 | 210 | 217 | 198 | 192.5 | 208 | 201 | 209 | 220 | 240 | 224 | 209 | 201 | 201.9 | 201.9668 | 193 | 207 | 190 | 230 | 208 | 200 | 225 | 213 | 208 | 205 | 202.8 | 223.4 | 204.1679 |  |  |  | 1
| 1 | 2 | 210.48 | 205 | 212 | 198 | 189.5 | 205 | 219 | 211 | 227 | 223 | 224 | 197 | 204 | 203 | 206.8686 | 195 | 209 | 199 | 230 | 203 | 188 | 222 | 223 | 206 | 203 | 214.1333 | 220.293 | 205.2794 |  |  |  | #N/A
| 1 | 3 | 200.19 | 207 | 215 | 203 | 192 | 197 | 198 | 206 | 228 | 242 | 224 | 204 | 209 | 203.9 | 202.1348 | 197 | 209 | 192 | 220 | 207 | 199 | 212 | 224 | 205 | 196 | 198 | 223.977 | 205.8313 |  |  |  | 16
| 1 | 4 | 210.08 | 206 | 215 | 194 | 187.5 | 204 | 207 | 211 | 220 | 262 | 218 | 203 | 211 | 204.1 | 204.5118 | 197 | 207 | 196 | 230 | 199 | 191 | 223 | 228 | 206 | 204 | 205.6333 | 224.936 | 202.858 |  |  |  | 206.434
| 1 | 5 | 210.81 | 205 | 220 | 201 | 189.5 | 204 | 199 | 209 | 232 | 234 | 218 | 209 | 203 | 202.9 | 205.9348 | 196 | 201 | 198 | 230 | 195 | 198 | 220 | 226 | 204 | 205 | 200.4333 | 221.026 | 207.4236 |  |  |  | 85
| 1 | 6 | 214.11 | 205 | 214 | 200 | 192 | 210 | 208 | 203 | 223 | 233 | 224 | 207 | 210 | 206 | 204.9712 | 196 | 203 | 202 | 240 | 197 | 211 | 223 | 216 | 205 | 200 | 196.7333 | 219.745 | 201.2921
| Mean |  | 209.385 | 206.3333 | 215.5 | 199 | 190.5 | 204.6667 | 205.3333 | 208.1667 | 225 | 239 | 222 | 204.8333 | 206.3333 | 203.6333 | 204.398 | 195.6667 | 206 | 196.1667 | 230 | 201.5 | 197.8333 | 220.8333 | 221.6667 | 205.6667 | 202.1667 | 202.9556 | 222.2295 | 204.4754
| Median |  | 210.56 | 205.5 | 215 | 199 | 190.75 | 204.5 | 204 | 209 | 225 | 237 | 224 | 205.5 | 206.5 | 203.45 | 204.7415 | 196 | 207 | 197 | 230 | 201 | 198.5 | 222.5 | 223.5 | 205.5 | 203.5 | 201.6167 | 222.213 | 204.7236
| Std Dev. |  | 4.7363 | 1.9664 | 2.7386 | 3.0984 | 1.9748 | 4.4572 | 7.8655 | 3.1252 | 4.8166 | 13.0843 | 3.0984 | 4.5789 | 4.1793 | 1.4024 | 1.992 | 1.5055 | 3.2863 | 4.4907 | 6.3246 | 5.3572 | 8.0353 | 4.6224 | 5.8878 | 1.3663 | 3.5449 | 6.3525 | 2.1504 | 2.1918
| Rel.Std.Dev. |  | 2.262% | 0.953% | 1.271% | 1.557% | 1.037% | 2.178% | 3.831% | 1.501% | 2.141% | 5.475% | 1.396% | 2.235% | 2.026% | 0.6887% | 0.9746% | 0.7694% | 1.595% | 2.289% | 2.75% | 2.659% | 4.062% | 2.093% | 2.656% | 0.6643% | 1.753% | 3.13% | 0.9676% | 1.072%
| PDM3 |  | 1.43% | -0.04875% | 4.392% | -3.601% | -7.719% | -0.8561% | -0.5332% | 0.8393% | 8.994% | 15.78% | 7.54% | -0.7754% | -0.04875% | -1.357% | -0.9863% | -5.216% | -0.2102% | -4.974% | 11.42% | -2.39% | -4.166% | 6.975% | 7.379% | -0.3717% | -2.067% | -1.685% | 7.652% | -0.9488%
| Z-Score (Absolute) |  | 0.66 | 0.14 | 1.71 | 1.11 | 2.57 | 0.14 | 0.03 | 0.46 | 3.34 | 5.73 | 2.82 | 0.11 | 0.14 | 0.32 | 0.19 | 1.68 | 0.09 | 1.6 | 4.19 | 0.68 | 1.31 | 2.62 | 2.77 | 0.03 | 0.57 | 0.44 | 2.86 | 0.18
| Analytical results for Be in OREAS 45f (Certified Value 1.2 ppm) | [32 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*OES |  | 2
2
| 1 | 1 | 1.0504 | 1.3 | 1.3 | 1.2 | 4.0682 | 1.3 | 1.26 | 1.16 | 1.1 | < 5 | 1.3 | 1 | 1 | 1.2 | 1.37 | 1.227 | 1.07 | 1.3 | 1.12 | 1.44 | 1.2 | 1.31 | 1.1 | 1.1 | 1.18 | 1.29 | 1.9 | 1.19 | 0.94 | 1.2465 |  | 1
| 1 | 2 | 1.0692 | 1.4 | 1.1 | 1.3 | 3.8632 | 1.3 | 0.85 | 1.18 | 1.1 | < 5 | 1.2 | 1.5 | 1 | 1.2 | 1.18 | 1.149 | 1.03 | 1.3 | 1.53 | 1.62 | 1.13 | 1.19 | 1 | 1.1 | 1.34 | 1.17 | 1.9 | 1.2433 | 0.905 | 1.2916 |  | #N/A
| 1 | 3 | 1.0491 | 1.2 | 1.3 | 1.2 | 3.8247 | 1.2 | 1.15 | 0.99 | 1.1 | < 5 | 1.3 | 1 | 1 | 1.2 | 1.32 | 1.2286 | 0.813 | 1.4 | 1.7 | 1.44 | 0.98 | 1.18 | 1 | 1.2 | 1.49 | 1.12 | 1.9 | 1.2133 | 0.908 | 1.2772 |  | 16
| 1 | 4 | 1.0906 | 1.4 | 1.3 | 1.2 | 3.7068 | 1.3 | 0.88 | 1.17 | 1.1 | < 5 | 1.2 | 1.5 | 1 | 1.2 | 1.23 | 1.1584 | 1.06 | 1.3 | 1.53 | 1.26 | 1.29 | 1.28 | 1 | 1.3 | 1.85 | 1.1 | 1.9 | 1.18 | 0.919 | 1.2881 |  | 1.2048
| 1 | 5 | 1.0597 | 1.2 | 1.3 | 1.2 | 3.8927 | 1.1 | 1.28 | 1.12 | 1.1 | < 5 | 1.2 | 1.5 | 2 | 1.2 | 1.26 | 1.1905 | 0.824 | 1.3 | 1.11 | 1.37 | 1.03 | 1.11 | 1 | 1.2 | 1.68 | 1.14 | 1.9 | 1.2267 | 0.906 | 1.2829 |  | 86
| 1 | 6 | 1.0307 | 1.1 | 1.4 | 1.3 | 3.6253 | 1.2 | 1.17 | 1.14 | 1.1 | < 5 | 1.3 | 1 | < 1 | 1.2 | 1.29 | 1.2407 | 0.828 | 1.4 | 1.24 | 1.14 | 1.12 | 1.19 | 1 | 1.2 | 1.4 | 1.1 | 1.9 | 1.1867 | 0.896 | 1.2745
| Mean |  | 1.0583 | 1.2667 | 1.2833 | 1.2333 | 3.8302 | 1.2333 | 1.0983 | 1.1267 | 1.1 |  | 1.25 | 1.25 | 1.2 | 1.2 | 1.275 | 1.199 | 0.9375 | 1.3333 | 1.3717 | 1.3783 | 1.125 | 1.21 | 1.0167 | 1.1833 | 1.49 | 1.1533 | 1.9 | 1.2067 | 0.91233 | 1.2768
| Median |  | 1.055 | 1.25 | 1.3 | 1.2 | 3.8439 | 1.25 | 1.16 | 1.15 | 1.1 |  | 1.25 | 1.25 | 1 | 1.2 | 1.275 | 1.2087 | 0.929 | 1.3 | 1.385 | 1.405 | 1.125 | 1.19 | 1 | 1.2 | 1.445 | 1.13 | 1.9 | 1.2017 | 0.907 | 1.28
| Std Dev. |  | 0.020358 | 0.12111 | 0.098319 | 0.05164 | 0.15419 | 0.08165 | 0.18777 | 0.070333 | 2.4324e-16 |  | 0.054772 | 0.27386 | 0.44721 | 0 | 0.067157 | 0.039032 | 0.12766 | 0.05164 | 0.24782 | 0.16546 | 0.11221 | 0.072938 | 0.040825 | 0.075277 | 0.24183 | 0.072019 | 2.4324e-16 | 0.02521 | 0.015423 | 0.016164
| Rel.Std.Dev. |  | 1.924% | 9.561% | 7.661% | 4.187% | 4.026% | 6.62% | 17.1% | 6.243% | 2.211e-14% |  | 4.382% | 21.91% | 37.27% | 0% | 5.267% | 3.255% | 13.62% | 3.873% | 18.07% | 12% | 9.974% | 6.028% | 4.016% | 6.361% | 16.23% | 6.244% | 1.28e-14% | 2.089% | 1.69% | 1.266%
| PDM3 |  | -12.16% | 5.139% | 6.522% | 2.372% | 217.9% | 2.372% | -8.834% | -6.482% | -8.695% |  | 3.755% | 3.755% | -0.3949% | -0.3949% | 5.83% | -0.4756% | -22.18% | 10.67% | 13.85% | 14.41% | -6.62% | 0.4351% | -15.61% | -1.778% | 23.68% | -4.268% | 57.71% | 0.1584% | -24.27% | 5.98%
| Z-Score (Absolute) |  | 1.18 | 0.22 | 0.34 | 0 | 17.51 | 0 | 0.91 | 0.72 | 0.9 | 8.54 | 0.11 | 0.11 | NA | 0.22 | 0.28 | 0.23 | 1.99 | 0.67 | 0.93 | 0.98 | 0.73 | 0.16 | 1.46 | 0.34 | 1.73 | 0.54 | 4.5 | 0.18 | 2.16 | 0.29
Results from laboratory 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 45f (Certified Value 0.21 ppm) | [32 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | Various [a] | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*MS |  |  | 2
|  |  | 21: see footer |  |  | 3
| 1 | 1 | < 0.25 | 0.37 | 0.21 | 0.22 | 0.35472 | 0.2 | 0.2 | 0.201 | 0.2 | < 5 | 0.2 | 0.1 | 0.3 | 0.2 | 0.21 | 0.20507 | 0.212 | 0.37 | 0.208 | 0.212 | 0.21 | 0.223 | 0.2 | 0.23 | 0.214 | 0.25 | 10 | 0.2 | 0.205 |  |  | 1
| 1 | 2 | < 0.25 | 0.33 | 0.21 | 0.23 | 0.32516 | 0.2 | 0.21 | 0.184 | 0.2 | < 5 | 0.2 | 0.1 | 0.3 | 0.2 | 0.21 | 0.21476 | 0.223 | 0.31 | 0.197 | 0.212 | 0.22 | 0.216 | 0.2 | 0.24 | 0.225 | 0.24 | 9 | 0.21667 | 0.18 |  |  | #N/A
| 1 | 3 | < 0.25 | 0.36 | 0.21 | 0.24 | 0.35681 | 0.2 | 0.21 | 0.188 | 0.2 | < 5 | 0.2 | 0.1 | 0.3 | < 5 | 0.22 | 0.19765 | 0.223 | 0.32 | 0.206 | 0.198 | 0.22 | 0.225 | 0.2 | 0.23 | 0.21 | 0.24 | 9 | 0.20333 | 0.191 |  |  | 16
| 1 | 4 | < 0.25 | 0.36 | 0.22 | 0.22 | 0.32011 | 0.2 | 0.19 | 0.196 | 0.2 | < 5 | 0.2 | 0.1 | 0.3 | 0.2 | 0.21 | 0.21036 | 0.208 | 0.36 | 0.193 | 0.201 | 0.24 | 0.222 | 0.2 | 0.24 | 0.221 | 0.24 | 9 | 0.2 | 0.176 |  |  | 0.21265
| 1 | 5 | < 0.25 | 0.4 | 0.22 | 0.22 | 0.39048 | 0.2 | 0.2 | 0.184 | 0.2 | < 5 | 0.2 | 0.1 | 0.3 | < 5 | 0.21 | 0.20082 | 0.22 | 0.38 | 0.196 | 0.214 | 0.22 | 0.222 | 0.2 | 0.24 | 0.219 | 0.24 | 8 | 0.19667 | 0.187 |  |  | 87
| 1 | 6 | < 0.25 | 0.32 | 0.21 | 0.25 | 0.38481 | 0.2 | 0.2 | 0.182 | 0.2 | < 5 | 0.2 | 0.1 | 0.3 | 0.2 | 0.2 | 0.20827 | 0.213 | 0.33 | 0.209 | 0.21 | 0.22 | 0.212 | 0.2 | 0.25 | 0.227 | 0.23 | 9 | 0.19667 | 0.186
| Mean |  |  | 0.35667 | 0.21333 | 0.23 | 0.35535 | 0.2 | 0.20167 | 0.18917 | 0.2 |  | 0.2 | 0.1 | 0.3 | 0.2 | 0.21 | 0.20615 | 0.2165 | 0.345 | 0.2015 | 0.20783 | 0.22167 | 0.22 | 0.2 | 0.23833 | 0.21933 | 0.24 | 9 | 0.20222 | 0.1875
| Median |  |  | 0.36 | 0.21 | 0.225 | 0.35577 | 0.2 | 0.2 | 0.186 | 0.2 |  | 0.2 | 0.1 | 0.3 | 0.2 | 0.21 | 0.20667 | 0.2165 | 0.345 | 0.2015 | 0.211 | 0.22 | 0.222 | 0.2 | 0.24 | 0.22 | 0.24 | 9 | 0.2 | 0.1865
| Std Dev. |  |  | 0.028752 | 0.005164 | 0.012649 | 0.029182 | 3.0405e-17 | 0.0075277 | 0.0076529 | 3.0405e-17 |  | 3.0405e-17 | 1.5202e-17 | 0 | 0 | 0.0063246 | 0.0062946 | 0.0063482 | 0.02881 | 0.0069498 | 0.0066458 | 0.0098319 | 0.0049396 | 3.0405e-17 | 0.0075277 | 0.0064704 | 0.0063246 | 0.63246 | 0.0075031 | 0.010095
| Rel.Std.Dev. |  |  | 8.061% | 2.421% | 5.5% | 8.212% | 1.52e-14% | 3.733% | 4.046% | 1.52e-14% |  | 1.52e-14% | 1.52e-14% | 0% | 0% | 3.012% | 3.053% | 2.932% | 8.351% | 3.449% | 3.198% | 4.435% | 2.245% | 1.52e-14% | 3.158% | 2.95% | 2.635% | 7.027% | 3.71% | 5.384%
| PDM3 |  |  | 67.73% | 0.3237% | 8.161% | 67.11% | -5.947% | -5.163% | -11.04% | -5.947% |  | -5.947% | -52.97% | 41.08% | -5.947% | -1.244% | -3.053% | 1.813% | 62.24% | -5.241% | -2.263% | 4.243% | 3.459% | -5.947% | 12.08% | 3.145% | 12.86% | 4132% | -4.902% | -11.82%
| Z-Score (Absolute) |  | 4.16 | 6.37 | 0.14 | 0.61 | 6.31 | NA | 0.67 | 1.24 | NA | 103.81 | NA | NA | NA | NA | 0.3 | 0.47 | 0 | 5.84 | 0.68 | 0.39 | 0.23 | 0.16 | NA | 0.99 | 0.13 | 1.07 | NA | 0.65 | 1.32
Results from laboratories 12, 15, 18, 19, 20, 21, 33 & 37 were removed due to their 0.1 ppm or greater reading resolution.
[a]: Lab 21: Rnd 01 - 4A*MS, Rnd 01 - 4A*OES
| Analytical results for Ca in OREAS 45f (Certified Value 0.096 wt.%) | [32 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.11068 | 0.1 | 0.1 | 0.095 | 0.0887 | 0.11 | 0.11 | 0.09 | 0.1 | 0.0775 | 0.1 | 0.11 | 0.07 | 0.0956 | 0.090038 | 0.0906 | 0.09 | 0.09 | 0.11 | 0.1 | 0.102 | 0.09 | 0.1 | 0.1 | 0.1 | 0.093 | 0.093011 | 0.092129 |  |  |  | 1
| 1 | 2 | 0.1079 | 0.1 | 0.09 | 0.092 | 0.0889 | 0.11 | 0.11 | 0.09 | 0.1 | 0.0795 | 0.1 | 0.1 | 0.08 | 0.0942 | 0.092723 | 0.0897 | 0.09 | 0.09 | 0.12 | 0.1 | 0.104 | 0.1 | 0.1 | 0.1 | 0.1 | 0.089933 | 0.097878 | 0.091996 |  |  |  | #N/A
| 1 | 3 | 0.10955 | 0.09 | 0.09 | 0.091 | 0.0878 | 0.1 | 0.11 | 0.09 | 0.09 | 0.0785 | 0.1 | 0.11 | 0.08 | 0.0951 | 0.090859 | 0.0933 | 0.09 | 0.1 | 0.11 | 0.1 | 0.105 | 0.09 | 0.1 | 0.1 | 0.1 | 0.0901 | 0.096222 | 0.092567 |  |  |  | 16
| 1 | 4 | 0.10758 | 0.1 | 0.09 | 0.093 | 0.0872 | 0.11 | 0.12 | 0.09 | 0.1 | 0.0635 | 0.1 | 0.1 | 0.08 | 0.0944 | 0.090684 | 0.089 | 0.09 | 0.09 | 0.11 | 0.1 | 0.101 | 0.08 | 0.1 | 0.1 | 0.1 | 0.091167 | 0.0899 | 0.092137 |  |  |  | 0.095873
| 1 | 5 | 0.11049 | 0.11 | 0.09 | 0.091 | 0.0878 | 0.1 | 0.11 | 0.09 | 0.1 | 0.0645 | 0.1 | 0.11 | 0.08 | 0.0947 | 0.093119 | 0.0895 | 0.09 | 0.1 | 0.11 | 0.09 | 0.105 | 0.1 | 0.1 | 0.1 | 0.09 | 0.091967 | 0.091702 | 0.092283 |  |  |  | 88
| 1 | 6 | 0.11163 | 0.1 | 0.09 | 0.095 | 0.0894 | 0.11 | 0.1 | 0.09 | 0.09 | 0.07 | 0.1 | 0.09 | 0.08 | 0.0948 | 0.091128 | 0.0895 | 0.09 | 0.09 | 0.12 | 0.09 | 0.106 | 0.09 | 0.1 | 0.1 | 0.09 | 0.0924 | 0.095501 | 0.091289
| Mean |  | 0.10964 | 0.1 | 0.091667 | 0.092833 | 0.0883 | 0.10667 | 0.11 | 0.09 | 0.096667 | 0.07225 | 0.1 | 0.10333 | 0.078333 | 0.0948 | 0.091425 | 0.090267 | 0.09 | 0.093333 | 0.11333 | 0.096667 | 0.10383 | 0.091667 | 0.1 | 0.1 | 0.096667 | 0.091428 | 0.094036 | 0.092067
| Median |  | 0.11002 | 0.1 | 0.09 | 0.0925 | 0.08825 | 0.11 | 0.11 | 0.09 | 0.1 | 0.07375 | 0.1 | 0.105 | 0.08 | 0.09475 | 0.090994 | 0.0896 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1045 | 0.09 | 0.1 | 0.1 | 0.1 | 0.091567 | 0.094256 | 0.092133
| Std Dev. |  | 0.0016151 | 0.0063246 | 0.0040825 | 0.0018348 | 0.00082946 | 0.005164 | 0.0063246 | 1.5202e-17 | 0.005164 | 0.0072232 | 1.5202e-17 | 0.008165 | 0.0040825 | 0.000502 | 0.0012196 | 0.0015756 | 1.5202e-17 | 0.005164 | 0.005164 | 0.005164 | 0.0019408 | 0.0075277 | 1.5202e-17 | 1.5202e-17 | 0.005164 | 0.0012468 | 0.0030096 | 0.00042777
| Rel.Std.Dev. |  | 1.473% | 6.325% | 4.454% | 1.976% | 0.9394% | 4.841% | 5.75% | 1.689e-14% | 5.342% | 9.998% | 1.52e-14% | 7.902% | 5.212% | 0.5295% | 1.334% | 1.746% | 1.689e-14% | 5.533% | 4.556% | 5.342% | 1.869% | 8.212% | 1.52e-14% | 1.52e-14% | 5.342% | 1.364% | 3.201% | 0.4646%
| PDM3 |  | 14.36% | 4.305% | -4.387% | -3.17% | -7.899% | 11.26% | 14.74% | -6.125% | 0.8283% | -24.64% | 4.305% | 7.782% | -18.29% | -1.119% | -4.639% | -5.847% | -6.125% | -2.649% | 18.21% | 0.8283% | 8.303% | -4.387% | 4.305% | 4.305% | 0.8283% | -4.636% | -1.916% | -3.97%
| Z-Score (Absolute) |  | 2.32 | 0.85 | 0.42 | 0.24 | 0.93 | 1.87 | 2.38 | 0.67 | 0.34 | 3.38 | 0.85 | 1.36 | 2.46 | 0.06 | 0.46 | 0.63 | 0.67 | 0.17 | 2.89 | 0.34 | 1.44 | 0.42 | 0.85 | 0.85 | 0.34 | 0.46 | 0.06 | 0.36
| Analytical results for Cd in OREAS 45f (Indicative Value 0.024 ppm) | [32 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  | 3
3
| 1 | 1 | < 0.25 | 0.02 | < 0.02 | 0.05 | 3.5091 | < 0.5 | < 0.01 | 0.01 | < 0.1 | < 0.5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.02 | 0.020915 | < 0.02 | 0.02 | 0.007 | 0.013 | < 0.02 | < 0.005 | < 1 | 0.03 | 0.005 | < 0.02 | < 1 | 0.02 | 0.21 | 0.0508 |  | 1
| 1 | 2 | < 0.25 | 0.03 | < 0.02 | 0.05 | 3.2667 | < 0.5 | 0.02 | 0.014 | < 0.1 | < 0.5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.02 | < 0.02 | < 0.02 | 0.02 | 0.01 | 0.018 | 0.03 | < 0.005 | < 1 | < 0.02 | 0.013 | < 0.02 | < 1 | 0.035 | 0.188 | 0.058 |  | 10
| 1 | 3 | < 0.25 | 0.02 | < 0.02 | 0.06 | 3.586 | < 0.5 | 0.04 | 0.008 | < 0.1 | < 0.5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.02 | < 0.02 | < 0.02 | 0.02 | 0.017 | < 0.005 | < 0.02 | < 0.005 | < 1 | < 0.02 | 0.016 | < 0.02 | < 1 | 0.025 | 0.23 | 0.0633 |  | 16
| 1 | 4 | < 0.25 | 0.03 | < 0.02 | 0.05 | 3.6161 | < 0.5 | < 0.01 | < 0.005 | < 0.1 | < 0.5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.02 | < 0.02 | < 0.02 | 0.02 | < 0.005 | 0.01 | 0.02 | 0.006 | < 1 | < 0.02 | 0.014 | < 0.02 | < 1 | 0.02 | 0.15 | 0.0384 |  | 0.02433
| 1 | 5 | < 0.25 | 0.03 | < 0.02 | < 0.05 | 3.4875 | < 0.5 | < 0.01 | 0.005 | < 0.1 | < 0.5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | 0.03 | < 0.02 | < 0.02 | 0.02 | 0.008 | 0.007 | < 0.02 | < 0.005 | < 1 | < 0.02 | 0.015 | < 0.02 | < 1 | 0.023333 | 0.109 | 0.0433 |  | 23
| 1 | 6 | < 0.25 | 0.02 | < 0.02 | 0.06 | 3.5258 | < 0.5 | 0.03 | 0.005 | < 0.1 | < 0.5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | 0.04 | < 0.02 | < 0.02 | < 0.02 | < 0.005 | < 0.005 | < 0.02 | < 0.005 | < 1 | < 0.02 | 0.016 | < 0.02 | < 1 | 0.025 | 0.135 | 0.0379
| Mean |  | 04: 0.025; 07: 0.054; 10: 3.4985; 13: 0.03; 14: 0.0084; 23: 0.035; 24: 0.020915; 27: 0.02; 28: 0.0105; 29: 0.012; 30: 0.025; 31: 0.006; 34: 0.03; 35: 0.013167; 38: 0.024722; 39: 0.17033; 40: 0.048617
| Median |  | 04: 0.025; 07: 0.05; 10: 3.5174; 13: 0.03; 14: 0.008; 23: 0.035; 24: 0.020915; 27: 0.02; 28: 0.009; 29: 0.0115; 30: 0.025; 31: 0.006; 34: 0.03; 35: 0.0145; 38: 0.024167; 39: 0.169; 40: 0.04705
| Std Dev. |  | 04: 0.0054772; 07: 0.0054772; 10: 0.12346; 13: 0.01; 14: 0.0037815; 23: 0.0070711; 27: 0; 28: 0.0045092; 29: 0.0046904; 30: 0.0070711; 35: 0.0041673; 38: 0.0055193; 39: 0.046625; 40: 0.010544
| Rel.Std.Dev. |  | 04: 21.91%; 07: 10.14%; 10: 3.529%; 13: 33.33%; 14: 45.02%; 23: 20.2%; 27: 0%; 28: 42.95%; 29: 39.09%; 30: 28.28%; 35: 31.65%; 38: 22.33%; 39: 27.37%; 40: 21.69%
| PDM3 |  | 04: 2.753%; 07: 121.9%; 10: 14279%; 13: 23.3%; 14: -65.48%; 23: 43.85%; 24: -14.04%; 27: -17.8%; 28: -56.84%; 29: -50.68%; 30: 2.753%; 31: -75.34%; 34: 23.3%; 35: -45.88%; 38: 1.611%; 39: 600.1%; 40: 99.82%
| Z-Score (Absolute) |  | 3.91 | 0.01 | 0.57 | 0.95 | 135.36 | 8.78 | 0.28 | 0.67 | 0.98 | 8.78 | 0.98 | 8.78 | 0.98 | 0.98 | 0.25 | 0.5 | 0.57 | 0.25 | 0.66 | 0.62 | 0.38 | 0.84 | NA | 0.44 | 0.45 | 0.57 | 18.52 | 0 | 5.67 | 0.93
Results from laboratory 33 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Ce in OREAS 45f (Certified Value 28.8 ppm) | [32 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 15 | 18 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*OES | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 33.4458 | 26.57 | 29.9 | 47.725 | 29.3 | 29.11 | 23.7 | 30 | 43 | 32 | 30 | 27.99 | 29.6054 | 26.1 | 11.7627 | 21.9 | 30.4 | 27.3 | 35.5 | 28.3 | 32.2 | 27.02 | 27.6 | 33.7 | 27.5433 | 26.783 | 26.3762 |  |  |  |  | 1
| 1 | 2 | 33.2183 | 27.89 | 29.5 | 43.0262 | 30.1 | 30.04 | 26.8 | 30 | 43.6 | 28 | 33 | 28.17 | 29.8265 | 25.6 | 11.0487 | 24 | 28.4 | 27.7 | 29.3 | 28.3 | 35.3 | 29.3 | 27.5 | 32.4 | 28.9633 | 26.617 | 27.6819 |  |  |  |  | #N/A
| 1 | 3 | 33.3183 | 25.97 | 28.9 | 40.0327 | 29.4 | 29.75 | 23.7 | 30 | 42.9 | 29 | 29 | 28.32 | 29.3684 | 25.7 | 13.8607 | 21.8 | 28.7 | 26.9 | 29.7 | 28.2 | 30.8 | 25.2 | 28 | 32.8 | 27.2467 | 26.591 | 26.5926 |  |  |  |  | 16
| 1 | 4 | 33.2296 | 27.47 | 29.7 | 49.9274 | 29.7 | 27.91 | 27 | 30 | 36.4 | 30 | 34 | 28.82 | 29.3408 | 26.1 | 10.2638 | 21.3 | 26.4 | 27.7 | 30.1 | 28.1 | 35 | 25.2 | 27.8 | 31.3 | 28.1033 | 26.56 | 24.9531 |  |  |  |  | 28.8195
| 1 | 5 | 33.2234 | 26.81 | 30.3 | 41.7955 | 29.2 | 29.93 | 21.2 | 29 | 34.6 | 32 | 31 | 28.63 | 29.8042 | 26 | 9.9511 | 21.2 | 26.7 | 27.8 | 29.9 | 26.3 | 30.7 | 28.14 | 28.1 | 32.2 | 27.4267 | 26.523 | 27.5145 |  |  |  |  | 89
| 1 | 6 | 33.1653 | 27.48 | 27.7 | 35.6448 | 29.2 | 31.45 | 18.15 | 30 | 33.1 | 31 | 31 | 29.14 | 29.6925 | 26 | 12.1395 | 19.4 | 27.3 | 28.2 | 29.8 | 26.4 | 32 | 28.82 | 27.3 | 32 | 27.6767 | 26.432 | 26.8647
| Mean |  | 33.2668 | 27.0317 | 29.3333 | 43.0253 | 29.4833 | 29.6983 | 23.425 | 29.8333 | 38.9333 | 30.3333 | 31.3333 | 28.5117 | 29.6063 | 25.9167 | 11.5044 | 21.6 | 27.9833 | 27.6 | 30.7167 | 27.6 | 32.6667 | 27.28 | 27.7167 | 32.4 | 27.8267 | 26.5843 | 26.6638
| Median |  | 33.2265 | 27.14 | 29.6 | 42.4109 | 29.35 | 29.84 | 23.7 | 30 | 39.65 | 30.5 | 31 | 28.475 | 29.649 | 26 | 11.4057 | 21.55 | 27.85 | 27.7 | 29.85 | 28.15 | 32.1 | 27.58 | 27.7 | 32.3 | 27.61 | 26.5755 | 26.7286
| Std Dev. |  | 0.10061 | 0.70983 | 0.92448 | 5.1897 | 0.35449 | 1.165 | 3.3793 | 0.40825 | 4.7597 | 1.633 | 1.8619 | 0.43116 | 0.21081 | 0.2137 | 1.427 | 1.4819 | 1.4932 | 0.44721 | 2.3583 | 0.9716 | 2.0196 | 1.7839 | 0.30605 | 0.80747 | 0.62733 | 0.11674 | 0.98131
| Rel.Std.Dev. |  | 0.3024% | 2.626% | 3.152% | 12.06% | 1.202% | 3.923% | 14.43% | 1.368% | 12.23% | 5.383% | 5.942% | 1.512% | 0.712% | 0.8246% | 12.4% | 6.861% | 5.336% | 1.62% | 7.678% | 3.52% | 6.182% | 6.539% | 1.104% | 2.492% | 2.254% | 0.4391% | 3.68%
| PDM3 |  | 15.43% | -6.204% | 1.783% | 49.29% | 2.303% | 3.049% | -18.72% | 3.518% | 35.09% | 5.253% | 8.723% | -1.068% | 2.73% | -10.07% | -60.08% | -25.05% | -2.901% | -4.232% | 6.583% | -4.232% | 13.35% | -5.342% | -3.827% | 12.42% | -3.445% | -7.756% | -7.48%
| Z-Score (Absolute) |  | 1.76 | 0.55 | 0.3 | 5.37 | 0.36 | 0.44 | 1.88 | 0.49 | 3.86 | 0.67 | 1.04 | 0 | 0.41 | 0.96 | 6.3 | 2.56 | 0.2 | 0.34 | 0.82 | 0.34 | 1.54 | 0.46 | 0.29 | 1.44 | 0.25 | 0.71 | 0.68
| Analytical results for Co in OREAS 45f (Certified Value 44.5 ppm) | [32 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS | 1
2
| 1 | 1 | 43.3317 | 43.9 | 45.1 | 46 | 41.8967 | 45 | 43.5 | 44.3 | 47.6 | 46 | 47 | 45 | 47.4 | 45.2 | 43.5 | 44.0845 | 42.3 | 43.6534 | 40 | 41.8 | 42.3 | 45.8 | 45.3 | 45.8 | 50.2 | 45.6 | 52.9 | 43 | 43.3667 | 54.846 | 44.5793 | 1
| 1 | 2 | 43.1396 | 45.8 | 44.6 | 49 | 41.6524 | 40 | 45.6 | 44 | 40.7 | 47 | 45 | 44 | 46.5 | 52 | 43.8 | 46.0437 | 43.5 | 44.0562 | 41.1 | 41.7 | 44.4 | 44.4 | 44.3 | 44 | 49.3 | 45.5 | 51.5 | 42 | 45.9333 | 54.406 | 44.4136 | 3
| 1 | 3 | 43.3777 | 45.3 | 46.1 | 47 | 40.3471 | 45 | 44.1 | 42.5 | 40.8 | 47 | 47 | 44 | 47 | 44.1 | 43.7 | 44.8771 | 43.9 | 43.8619 | 41.2 | 41.9 | 42.9 | 44.3 | 44.7 | 42.8 | 48.2 | 45.9 | 49.9 | 42 | 43.2667 | 55.285 | 44.9087 | 16
| 1 | 4 | 43.2209 | 44.9 | 46.1 | 45 | 42.8845 | 45 | 43.5 | 43.7 | 40.2 | 46 | 48 | 45 | 48.8 | 48.1 | 44.6 | 44.7174 | 43.1 | 43.2013 | 39.8 | 40.8 | 42.1 | 45.5 | 43.9 | 43.2 | 50.2 | 45.3 | 48.6 | 44 | 44.1667 | 55.317 | 44.8805 | 44.5486
| 1 | 5 | 43.3758 | 45.7 | 46.1 | 47 | 42.7237 | 45 | 43.7 | 42.7 | 39.8 | 46 | 47 | 47 | 48.7 | 43.9 | 44 | 45.5546 | 42.8 | 46.3939 | 42 | 40.7 | 44 | 44.2 | 42.3 | 46.9 | 50.5 | 44.9 | 48.6 | 43 | 43.4667 | 54.85 | 45.2641 | 90
| 1 | 6 | 43.217 | 43.2 | 44.9 | 47 | 42.2923 | 45 | 46 | 41.3 | 39.4 | 47 | 46 | 43 | 48.2 | 45.6 | 45.2 | 46.2545 | 41.7 | 44.6521 | 40.5 | 42.9 | 43.5 | 44.3 | 42.9 | 46.8 | 51.1 | 45.1 | 46.5 | 42 | 42.6667 | 54.778 | 44.5032
| Mean |  | 43.2771 | 44.8 | 45.4833 | 46.8333 | 41.9661 | 44.1667 | 44.4 | 43.0833 | 41.4167 | 46.5 | 46.6667 | 44.6667 | 47.7667 | 46.4833 | 44.1333 | 45.2553 | 42.8833 | 44.3031 | 40.7667 | 41.6333 | 43.2 | 44.75 | 43.9 | 44.9167 | 49.9167 | 45.3833 | 49.6667 | 42.6667 | 43.8111 | 54.9137 | 44.7582
| Median |  | 43.2763 | 45.1 | 45.6 | 47 | 42.0945 | 45 | 43.9 | 43.2 | 40.45 | 46.5 | 47 | 44.5 | 47.8 | 45.4 | 43.9 | 45.2159 | 42.95 | 43.959 | 40.8 | 41.75 | 43.2 | 44.35 | 44.1 | 44.9 | 50.2 | 45.4 | 49.25 | 42.5 | 43.4167 | 54.848 | 44.7299
| Std Dev. |  | 0.098504 | 1.0431 | 0.69402 | 1.3292 | 0.92178 | 2.0412 | 1.1136 | 1.1286 | 3.0753 | 0.54772 | 1.0328 | 1.3663 | 0.94375 | 3.0928 | 0.64395 | 0.83817 | 0.80104 | 1.13 | 0.82624 | 0.80911 | 0.9252 | 0.7064 | 1.1243 | 1.8181 | 1.0226 | 0.36009 | 2.2879 | 0.8165 | 1.1448 | 0.34215 | 0.31916
| Rel.Std.Dev. |  | 0.2276% | 2.328% | 1.526% | 2.838% | 2.196% | 4.622% | 2.508% | 2.62% | 7.425% | 1.178% | 2.213% | 3.059% | 1.976% | 6.654% | 1.459% | 1.852% | 1.868% | 2.551% | 2.027% | 1.943% | 2.142% | 1.579% | 2.561% | 4.048% | 2.049% | 0.7934% | 4.607% | 1.914% | 2.613% | 0.6231% | 0.7131%
| PDM3 |  | -2.854% | 0.5643% | 2.098% | 5.129% | -5.797% | -0.8574% | -0.3336% | -3.289% | -7.03% | 4.38% | 4.754% | 0.265% | 7.224% | 4.343% | -0.9322% | 1.586% | -3.738% | -0.5511% | -8.489% | -6.544% | -3.027% | 0.4521% | -1.456% | 0.8262% | 12.05% | 1.874% | 11.49% | -4.224% | -1.656% | 23.27% | 0.4705%
| Z-Score (Absolute) |  | 0.64 | 0.06 | 0.38 | 1 | 1.24 | 0.23 | 0.12 | 0.73 | 1.49 | 0.84 | 0.92 | 0 | 1.43 | 0.84 | 0.25 | 0.27 | 0.82 | 0.17 | 1.79 | 1.39 | 0.67 | 0.04 | 0.35 | 0.11 | 2.41 | 0.33 | 2.3 | 0.92 | 0.39 | 4.71 | 0.04
| Analytical results for Cr in OREAS 45f (Certified Value 417 ppm) | [32 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 391 | 385 | 418 | 382.5 | 370 | 490 | 428 | 585 | 377 | 410 | 400 | 481 | 443 | 439 | 428.9853 | 384 | 405 | 435 | 424 | 489 | 459 | 350 | 398 | 453 | 410 | 393 | 404.3333 | 395 | 439.0508 |  |  | 1
| 1 | 2 | 397 | 410 | 406 | 374 | 380 | 503 | 425 | 589 | 346 | 395 | 400 | 446 | 449 | 445 | 446.3417 | 394 | 416 | 447 | 423 | 498 | 443 | 370 | 399 | 464 | 382 | 383 | 397.6667 | 396.982 | 445.6755 |  |  | #N/A
| 1 | 3 | 405 | 370 | 402 | 377.5 | 380 | 459 | 416 | 592 | 392 | 420 | 380 | 465 | 431 | 418 | 431.9706 | 395 | 433 | 450 | 432 | 477 | 457 | 355 | 370 | 451 | 396 | 387 | 405 | 410.139 | 450.7653 |  |  | 16
| 1 | 4 | 409 | 436 | 423 | 371.5 | 400 | 512 | 427 | 577 | 377 | 430 | 410 | 467 | 441 | 463 | 431.5282 | 390 | 414 | 439 | 433 | 493 | 444 | 350 | 403 | 459 | 408 | 394 | 390 | 397.965 | 443.6604 |  |  | 417.3037
| 1 | 5 | 407 | 388 | 406 | 371 | 380 | 458 | 424 | 573 | 368 | 405 | 400 | 472 | 437 | 448 | 434.7848 | 392 | 413 | 443 | 433 | 481 | 432 | 370 | 398 | 452 | 414 | 385 | 360.6667 | 412.912 | 439.5107 |  |  | 91
| 1 | 6 | 402 | 419 | 416 | 386.5 | 360 | 434 | 420 | 579 | 353 | 410 | 340 | 472 | 453 | 419 | 442.1178 | 387 | 419 | 438 | 441 | 499 | 446 | 360 | 409 | 446 | 405 | 383 | 354.3333 | 401.308 | 440.9257
| Mean |  | 401.8333 | 401.3333 | 411.8333 | 377.1667 | 378.3333 | 476 | 423.3333 | 582.5 | 368.8333 | 411.6667 | 388.3333 | 467.1667 | 442.3333 | 438.6667 | 435.9547 | 390.3333 | 416.6667 | 442 | 431 | 489.5 | 446.8333 | 359.1667 | 396.1667 | 454.1667 | 402.5 | 387.5 | 385.3333 | 402.3843 | 443.2647
| Median |  | 403.5 | 399 | 411 | 375.75 | 380 | 474.5 | 424.5 | 582 | 372.5 | 410 | 400 | 469.5 | 442 | 442 | 433.3777 | 391 | 415 | 441 | 432.5 | 491 | 445 | 357.5 | 398.5 | 452.5 | 406.5 | 386 | 393.8333 | 399.6365 | 442.293
| Std Dev. |  | 6.7651 | 24.5574 | 8.3046 | 6.2583 | 13.2916 | 30.3249 | 4.5461 | 7.3689 | 16.9873 | 12.1106 | 25.6255 | 11.7544 | 7.9666 | 17.5119 | 6.8012 | 4.2269 | 9.2664 | 5.7271 | 6.6633 | 8.9833 | 9.9482 | 9.1742 | 13.4969 | 6.3692 | 11.726 | 4.8888 | 22.3209 | 7.4212 | 4.462
| Rel.Std.Dev. |  | 1.684% | 6.119% | 2.016% | 1.659% | 3.513% | 6.371% | 1.074% | 1.265% | 4.606% | 2.942% | 6.599% | 2.516% | 1.801% | 3.992% | 1.56% | 1.083% | 2.224% | 1.296% | 1.546% | 1.835% | 2.226% | 2.554% | 3.407% | 1.402% | 2.913% | 1.262% | 5.793% | 1.844% | 1.007%
| PDM3 |  | -3.707% | -3.827% | -1.311% | -9.618% | -9.339% | 14.07% | 1.445% | 39.59% | -11.62% | -1.351% | -6.942% | 11.95% | 5.998% | 5.119% | 4.469% | -6.463% | -0.1527% | 5.918% | 3.282% | 17.3% | 7.076% | -13.93% | -5.065% | 8.834% | -3.547% | -7.142% | -7.661% | -3.575% | 6.221%
| Z-Score (Absolute) |  | 0.25 | 0.27 | 0 | 0.88 | 0.85 | 1.63 | 0.29 | 4.34 | 1.09 | 0 | 0.6 | 1.41 | 0.78 | 0.68 | 0.61 | 0.55 | 0.12 | 0.77 | 0.49 | 1.98 | 0.89 | 1.34 | 0.4 | 1.08 | 0.24 | 0.62 | 0.67 | 0.24 | 0.8
| Analytical results for Cs in OREAS 45f (Certified Value 3.65 ppm) | [32 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 4.1419 | 3.99 | 4 | 3.7 | 8.483 | 3.8 | 3.71 | 3.55 | 3.4 | 4.42 | 3.5 | 3.62 | 3.5393 | 3.66 | 3.43 | 3.98 | 3.5 | 4.08 | 3.53 | 3.85 | < 5 | 3.47 | 3.36 | 3.55 | 3.575 | 3.6635 | [5 empty] | 1
| 1 | 2 | 4.1471 | 3.94 | 4 | 3.8 | 8.1691 | 3.7 | 3.67 | 3.59 | 3.3 | 4.25 | 3.4 | 3.75 | 3.8078 | 3.53 | 3.32 | 3.69 | 3.64 | 3.97 | 3.48 | 4.29 | < 5 | 3.51 | 3.51 | 3.8 | 3.254 | 3.6309 | [5 empty] | #N/A
| 1 | 3 | 4.0986 | 3.36 | 4 | 3.7 | 8.7675 | 3.5 | 3.73 | 3.46 | 3.8 | 4.54 | 3.3 | 3.86 | 3.6837 | 3.6 | 3.38 | 3.8 | 3.5 | 3.93 | 3.53 | 3.81 | < 5 | 3.6 | 3.35 | 3.4867 | 3.458 | 3.6967 | [5 empty] | 16
| 1 | 4 | 4.1143 | 3.95 | 4 | 3.5 | 8.1316 | 3.6 | 3.7 | 3.63 | 3.5 | 3.47 | 3.5 | 3.85 | 3.6842 | 3.68 | 3.18 | 3.75 | 3.58 | 4.16 | 3.41 | 4.12 | < 5 | 3.84 | 3.32 | 3.6067 | 3.123 | 3.6953 | [5 empty] | 3.6493
| 1 | 5 | 4.0535 | 3.73 | 4 | 3.7 | 8.0155 | 3.6 | 3.58 | 3.58 | 3.3 | 3.71 | 3.8 | 3.76 | 3.7467 | 3.68 | 3.16 | 3.79 | 3.5 | 3.76 | 3.44 | 3.73 | < 5 | 3.73 | 3.35 | 3.5467 | 3.324 | 3.6586 | [5 empty] | 92
| 1 | 6 | 4.0511 | 3.75 | 4 | 3.7 | 8.5341 | 3.7 | 3.72 | 3.51 | 3.5 | 3.65 | 3.3 | 3.61 | 3.9134 | 3.58 | 3.1 | 3.82 | 3.61 | 3.86 | 3.38 | 3.98 | < 5 | 3.56 | 3.2 | 3.46 | 3.433 | 3.6341
| Mean |  | 4.1011 | 3.7867 | 4 | 3.6833 | 8.3501 | 3.65 | 3.685 | 3.5533 | 3.4667 | 4.0067 | 3.4667 | 3.7417 | 3.7292 | 3.6217 | 3.2617 | 3.805 | 3.555 | 3.96 | 3.4617 | 3.9633 |  | 3.6183 | 3.3483 | 3.575 | 3.3612 | 3.6632
| Median |  | 4.1064 | 3.845 | 4 | 3.7 | 8.326 | 3.65 | 3.705 | 3.565 | 3.45 | 3.98 | 3.45 | 3.755 | 3.7155 | 3.63 | 3.25 | 3.795 | 3.54 | 3.95 | 3.46 | 3.915 |  | 3.58 | 3.35 | 3.5483 | 3.3785 | 3.661
| Std Dev. |  | 0.041778 | 0.23585 | 0 | 0.098319 | 0.28922 | 0.10488 | 0.055408 | 0.060882 | 0.18619 | 0.45116 | 0.18619 | 0.10797 | 0.12692 | 0.061455 | 0.13333 | 0.097314 | 0.063166 | 0.14519 | 0.062423 | 0.21087 |  | 0.14077 | 0.099079 | 0.12172 | 0.16115 | 0.02851
| Rel.Std.Dev. |  | 1.019% | 6.229% | 0% | 2.669% | 3.464% | 2.873% | 1.504% | 1.713% | 5.371% | 11.26% | 5.371% | 2.886% | 3.403% | 1.697% | 4.088% | 2.558% | 1.777% | 3.666% | 1.803% | 5.321% |  | 3.891% | 2.959% | 3.405% | 4.795% | 0.7783%
| PDM3 |  | 12.38% | 3.764% | 9.61% | 0.9325% | 128.8% | 0.01912% | 0.9782% | -2.63% | -5.005% | 9.793% | -5.005% | 2.531% | 2.189% | -0.7573% | -10.62% | 4.267% | -2.584% | 8.514% | -5.142% | 8.605% |  | -0.8486% | -8.247% | -2.036% | -7.896% | 0.3804%
| Z-Score (Absolute) |  | 1.96 | 0.59 | NA | 0.15 | 20.45 | 0 | 0.15 | 0.42 | 0.8 | 1.55 | 0.8 | 0.4 | 0.34 | 0.12 | 1.69 | 0.67 | 0.41 | 1.35 | 0.82 | 1.36 | 5 | 0.14 | 1.31 | 0.33 | 1.26 | 0.06
Results from laboratory 5 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 45f (Certified Value 363 ppm) | [32 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*MS | 0
0
| 1 | 1 | 377.4744 | 357.7 | 347 | 353 | 352 | 378 | 380.9 | 348 | 364 | 342 | 380 | 378 | 390.1 | 355 | 364 | 353.6834 | 337 | 362.9008 | 382 | 346 | 353 | 371 | 370 | 369 | 395.8 | 364 | 374 | 363 | 334.4333 | 349.783 | 407.4747 | 1
| 1 | 2 | 372.7814 | 369.2 | 349 | 352 | 354 | 366 | 362.2 | 345 | 367 | 365 | 359 | 370 | 368.4 | 367 | 360 | 364.7008 | 336 | 366.4833 | 382 | 346 | 356 | 377 | 349 | 357 | 396.3 | 379 | 365 | 363 | 325.0333 | 348.001 | 409.5034 | #N/A
| 1 | 3 | 373.6635 | 364.6 | 343 | 348 | 350.5 | 360 | 382.6 | 337 | 365 | 359 | 389 | 388 | 389.5 | 341 | 361 | 356.9582 | 359 | 358.8183 | 391 | 353 | 362 | 362 | 369 | 359 | 379.4 | 373 | 370 | 347 | 326.7333 | 361.565 | 416.5797 | 16
| 1 | 4 | 372.2861 | 366.2 | 354 | 359 | 349.5 | 360 | 347.3 | 342 | 361 | 362 | 398 | 374 | 377.8 | 350 | 358 | 353.5796 | 340 | 358.2384 | 395 | 339 | 352 | 375 | 352 | 347 | 399.1 | 375 | 369 | 366 | 329.6333 | 355.298 | 412.3111 | 362.5909
| 1 | 5 | 372.7163 | 365.7 | 345 | 354 | 346.5 | 378 | 387.5 | 337 | 368 | 356 | 379 | 380 | 395.9 | 354 | 364 | 361.4208 | 336 | 373.1652 | 379 | 358 | 373 | 367 | 342 | 358 | 392.8 | 367 | 364 | 360 | 331.3667 | 353.36 | 414.6412 | 93
| 1 | 6 | 376.4522 | 363.1 | 340 | 355 | 355 | 370 | 385.1 | 333 | 367 | 361 | 375 | 372 | 380.8 | 357 | 369 | 365.5193 | 336 | 366.7055 | 376 | 352 | 361 | 381 | 353 | 374 | 393.3 | 365 | 361 | 350 | 327.7667 | 353.023 | 401.207
| Mean |  | 374.229 | 364.4167 | 346.3333 | 353.5 | 351.25 | 368.6667 | 374.2667 | 340.3333 | 365.3333 | 357.5 | 380 | 377 | 383.75 | 354 | 362.6667 | 359.3104 | 340.6667 | 364.3853 | 384.1667 | 349 | 359.5 | 372.1667 | 355.8333 | 360.6667 | 392.7833 | 370.5 | 367.1667 | 358.1667 | 329.1611 | 353.505 | 410.2862
| Median |  | 373.2224 | 365.15 | 346 | 353.5 | 351.25 | 368 | 381.75 | 339.5 | 366 | 360 | 379.5 | 376 | 385.15 | 354.5 | 362.5 | 359.1895 | 336.5 | 364.6921 | 382 | 349 | 358.5 | 373 | 352.5 | 358.5 | 394.55 | 370 | 367 | 361.5 | 328.7 | 353.1915 | 410.9072
| Std Dev. |  | 2.1888 | 3.8623 | 4.8854 | 3.6194 | 3.1105 | 8.165 | 15.9921 | 5.6451 | 2.582 | 8.167 | 13.2061 | 6.5422 | 10.0073 | 8.5323 | 3.8816 | 5.331 | 9.1141 | 5.6187 | 7.3052 | 6.6933 | 7.7653 | 6.9402 | 11.2679 | 9.5638 | 6.939 | 6.0581 | 4.7081 | 7.7825 | 3.3989 | 4.7451 | 5.543
| Rel.Std.Dev. |  | 0.5849% | 1.06% | 1.411% | 1.024% | 0.8855% | 2.215% | 4.273% | 1.659% | 0.7067% | 2.284% | 3.475% | 1.735% | 2.608% | 2.41% | 1.07% | 1.484% | 2.675% | 1.542% | 1.902% | 1.918% | 2.16% | 1.865% | 3.167% | 2.652% | 1.767% | 1.635% | 1.282% | 2.173% | 1.033% | 1.342% | 1.351%
| PDM3 |  | 3.21% | 0.5035% | -4.484% | -2.507% | -3.128% | 1.676% | 3.22% | -6.138% | 0.7563% | -1.404% | 4.801% | 3.974% | 5.836% | -2.369% | 0.02089% | -0.9048% | -6.047% | 0.4949% | 5.95% | -3.748% | -0.8525% | 2.641% | -1.864% | -0.5307% | 8.327% | 2.181% | 1.262% | -1.22% | -9.22% | -2.506% | 13.15%
| Z-Score (Absolute) |  | 0.85 | 0.13 | 1.2 | 0.67 | 0.84 | 0.44 | 0.85 | 1.64 | 0.2 | 0.38 | 1.28 | 1.05 | 1.55 | 0.64 | 0 | 0.25 | 1.62 | 0.13 | 1.58 | 1.01 | 0.23 | 0.7 | 0.5 | 0.15 | 2.22 | 0.58 | 0.33 | 0.33 | 2.46 | 0.67 | 3.5
| Analytical results for Dy in OREAS 45f (Certified Value 2.23 ppm) | [32 empty] | Certified
Dy | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab | 1
| No | No | 02: 02; 04: 12; 05: 13; 07: 15; 10: 18; 12: 23; 13: 24; 14: 27; 15: 33; 17: 38; 18: 39; 19: 40 | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*OES/MS; 15: 4A*MS; 17: 4A*MS; 18: 5A*OES; 19: 4A*MS | 2
3
| 1 | 1 | 02: 2.298; 04: 2.2; 05: 2.44; 07: 2.04; 10: 2.8; 12: 2.25; 13: 2.2541; 14: 2.27; 15: 1.9; 17: 2.15; 18: 0.337; 19: 2.6233 | 1
| 1 | 2 | 02: 2.2834; 04: 2.4; 05: 2.44; 07: 2.05; 10: 2.6; 12: 2.15; 13: 2.2692; 14: 2.31; 15: 2; 17: 2.2667; 18: 0.61; 19: 2.723 | #N/A
| 1 | 3 | 02: 2.2433; 04: 2.2; 05: 2.42; 07: 2.2; 10: 2.8; 12: 2.21; 13: 2.2843; 14: 2.28; 15: 1.9; 17: 2.1367; 18: 0.27; 19: 2.644 | 16
| 1 | 4 | 02: 2.3213; 04: 2.35; 05: 2.15; 07: 2.19; 10: 3; 12: 2.27; 13: 2.1882; 14: 2.21; 15: 2.1; 17: 2.18; 18: 0.342; 19: 2.5145 | 2.2254
| 1 | 5 | 02: 2.2654; 04: 2.25; 05: 2.42; 07: 1.94; 10: 2.6; 12: 2.31; 13: 2.2537; 14: 2.21; 15: 2.1; 17: 2.24; 18: 0.68; 19: 2.6458 | 94
| 1 | 6 | 02: 2.239; 04: 2.35; 05: 2.31; 07: 2.37; 10: 2.6; 12: 2.15; 13: 2.3399; 14: 2.25; 15: 1.9; 17: 2.21; 18: 0.577; 19: 2.5792
| Mean |  | 02: 2.2751; 04: 2.2917; 05: 2.3633; 07: 2.1317; 10: 2.7333; 12: 2.2233; 13: 2.2649; 14: 2.255; 15: 1.9833; 17: 2.1972; 18: 0.46933; 19: 2.6216
| Median |  | 02: 2.2744; 04: 2.3; 05: 2.42; 07: 2.12; 10: 2.7; 12: 2.23; 13: 2.2616; 14: 2.26; 15: 1.95; 17: 2.195; 18: 0.4595; 19: 2.6336
| Std Dev. |  | 02: 0.032051; 04: 0.08612; 05: 0.11535; 07: 0.15276; 10: 0.1633; 12: 0.06532; 13: 0.0493; 14: 0.039875; 15: 0.098319; 17: 0.051008; 18: 0.17275; 19: 0.0702
| Rel.Std.Dev. |  | 02: 1.409%; 04: 3.758%; 05: 4.881%; 07: 7.166%; 10: 5.974%; 12: 2.938%; 13: 2.177%; 14: 1.768%; 15: 4.957%; 17: 2.321%; 18: 36.81%; 19: 2.678%
| PDM3 |  | 02: 2.234%; 04: 2.98%; 05: 6.2%; 07: -4.21%; 10: 22.83%; 12: -0.0908%; 13: 1.777%; 14: 1.332%; 15: -10.88%; 17: -1.264%; 18: -78.91%; 19: 17.81%
| Z-Score (Absolute) |  | 02: 0.12; 04: 0.26; 05: 0.84; 07: 1.04; 10: 3.84; 12: 0.3; 13: 0.04; 14: 0.04; 15: 2.25; 17: 0.51; 18: 14.54; 19: 2.94
| Analytical results for Er in OREAS 45f (Certified Value 1.33 ppm) | [32 empty] | Certified
Er | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab | 1
| No | No | 02: 02; 04: 12; 05: 13; 07: 15; 10: 18; 12: 23; 13: 24; 14: 27; 15: 33; 17: 38; 18: 39; 19: 40 | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*OES/MS; 15: 4A*MS; 17: 4A*MS; 18: 5A*MS; 19: 4A*MS | 2
3
| 1 | 1 | 02: 1.2584; 04: 1.3; 05: 1.33; 07: 1.3; 10: 1.5; 12: 1.24; 13: 1.3431; 14: 1.28; 15: 1.3; 17: 1.2467; 18: 1.194; 19: 1.635 | 1
| 1 | 2 | 02: 1.2718; 04: 1.45; 05: 1.52; 07: 1.42; 10: 1.4; 12: 1.24; 13: 1.3005; 14: 1.3; 15: 1.2; 17: 1.3333; 18: 0.999; 19: 1.6659 | #N/A
| 1 | 3 | 02: 1.2545; 04: 1.25; 05: 1.4; 07: 1.42; 10: 1.5; 12: 1.27; 13: 1.3327; 14: 1.27; 15: 1.3; 17: 1.2633; 18: 1.052; 19: 1.6026 | 16
| 1 | 4 | 02: 1.2602; 04: 1.45; 05: 1.21; 07: 1.41; 10: 1.6; 12: 1.22; 13: 1.4516; 14: 1.29; 15: 1.3; 17: 1.2767; 18: 0.914; 19: 1.6178 | 1.3331
| 1 | 5 | 02: 1.2614; 04: 1.4; 05: 1.41; 07: 1.29; 10: 1.5; 12: 1.23; 13: 1.2606; 14: 1.24; 15: 1.4; 17: 1.2967; 18: 1.009; 19: 1.6528 | 95
| 1 | 6 | 02: 1.2563; 04: 1.4; 05: 1.35; 07: 1.37; 10: 1.4; 12: 1.21; 13: 1.4065; 14: 1.28; 15: 1.4; 17: 1.3633; 18: 1.031; 19: 1.5273
| Mean |  | 02: 1.2604; 04: 1.375; 05: 1.37; 07: 1.3683; 10: 1.4833; 12: 1.235; 13: 1.3492; 14: 1.2767; 15: 1.3167; 17: 1.2967; 18: 1.0332; 19: 1.6169
| Median |  | 02: 1.2593; 04: 1.4; 05: 1.375; 07: 1.39; 10: 1.5; 12: 1.235; 13: 1.3379; 14: 1.28; 15: 1.3; 17: 1.2867; 18: 1.02; 19: 1.6264
| Std Dev. |  | 02: 0.0061092; 04: 0.082158; 05: 0.10257; 07: 0.059805; 10: 0.075277; 12: 0.020736; 13: 0.069712; 14: 0.020656; 15: 0.075277; 17: 0.044272; 18: 0.091846; 19: 0.049499
| Rel.Std.Dev. |  | 02: 0.4847%; 04: 5.975%; 05: 7.487%; 07: 4.371%; 10: 5.075%; 12: 1.679%; 13: 5.167%; 14: 1.618%; 15: 5.717%; 17: 3.414%; 18: 8.89%; 19: 3.061%
| PDM3 |  | 02: -5.453%; 04: 3.141%; 05: 2.766%; 07: 2.641%; 10: 11.27%; 12: -7.361%; 13: 1.203%; 14: -4.235%; 15: -1.235%; 17: -2.735%; 18: -22.5%; 19: 21.29%
| Z-Score (Absolute) |  | 02: 0.99; 04: 0.58; 05: 0.51; 07: 0.49; 10: 2.06; 12: 1.34; 13: 0.22; 14: 0.77; 15: 0.22; 17: 0.5; 18: 4.11; 19: 3.89
| Analytical results for Eu in OREAS 45f (Certified Value 0.63 ppm) | [32 empty] | Certified
Eu | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab | 1
| No | No | 02: 02; 04: 12; 05: 13; 07: 15; 10: 18; 12: 23; 13: 24; 14: 27; 15: 33; 17: 38; 18: 39; 19: 40 | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*OES/MS; 15: 4A*MS; 17: 4A*MS; 18: 5A*MS; 19: 4A*MS | 2
2
| 1 | 1 | 02: 0.7337; 04: 0.6; 05: 0.65; 07: 0.48; 10: 0.96; 12: 0.62; 13: 0.63793; 14: 0.78; 15: 0.69; 17: 0.60333; 18: 0.599; 19: 0.6073 | 1
| 1 | 2 | 02: 0.7164; 04: 0.55; 05: 0.65; 07: 0.55; 10: 0.91; 12: 0.61; 13: 0.69217; 14: 0.79; 15: 0.54; 17: 0.64; 18: 0.532; 19: 0.6273 | #N/A
| 1 | 3 | 02: 0.7478; 04: 0.55; 05: 0.72; 07: 0.64; 10: 0.97; 12: 0.61; 13: 0.68625; 14: 0.77; 15: 0.67; 17: 0.60667; 18: 0.569; 19: 0.6166 | 16
| 1 | 4 | 02: 0.746; 04: 0.6; 05: 0.6; 07: 0.52; 10: 0.97; 12: 0.59; 13: 0.6456; 14: 0.76; 15: 0.66; 17: 0.61; 18: 0.509; 19: 0.568 | 0.63351
| 1 | 5 | 02: 0.7168; 04: 0.55; 05: 0.65; 07: 0.58; 10: 0.86; 12: 0.63; 13: 0.64819; 14: 0.77; 15: 0.74; 17: 0.59667; 18: 0.53; 19: 0.6402 | 96
| 1 | 6 | 02: 0.7244; 04: 0.55; 05: 0.65; 07: 0.57; 10: 0.84; 12: 0.59; 13: 0.66546; 14: 0.7; 15: 0.72; 17: 0.6; 18: 0.551; 19: 0.5873
| Mean |  | 02: 0.73085; 04: 0.56667; 05: 0.65333; 07: 0.55667; 10: 0.91833; 12: 0.60833; 13: 0.6626; 14: 0.76167; 15: 0.67; 17: 0.60944; 18: 0.54833; 19: 0.60778
| Median |  | 02: 0.72905; 04: 0.55; 05: 0.65; 07: 0.56; 10: 0.935; 12: 0.61; 13: 0.65683; 14: 0.77; 15: 0.68; 17: 0.605; 18: 0.5415; 19: 0.61195
| Std Dev. |  | 02: 0.013948; 04: 0.02582; 05: 0.038297; 07: 0.05465; 10: 0.057764; 12: 0.016021; 13: 0.022569; 14: 0.031885; 15: 0.070427; 17: 0.015694; 18: 0.032098; 19: 0.026503
| Rel.Std.Dev. |  | 02: 1.908%; 04: 4.556%; 05: 5.862%; 07: 9.817%; 10: 6.29%; 12: 2.634%; 13: 3.406%; 14: 4.186%; 15: 10.51%; 17: 2.575%; 18: 5.854%; 19: 4.361%
| PDM3 |  | 02: 15.37%; 04: -10.55%; 05: 3.13%; 07: -12.13%; 10: 44.96%; 12: -3.974%; 13: 4.593%; 14: 20.23%; 15: 5.761%; 17: -3.798%; 18: -13.44%; 19: -4.06%
| Z-Score (Absolute) |  | 02: 1.3; 04: 0.84; 05: 0.29; 07: 0.98; 10: 3.74; 12: 0.3; 13: 0.41; 14: 1.7; 15: 0.5; 17: 0.29; 18: 1.08; 19: 0.31
| Analytical results for Fe in OREAS 45f (Certified Value 14.65 wt.%) | [32 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 14.0663 | 14.67 | > 15 | 17.3 | 12.3287 | 15.2 | 15.01 | 14.2 | 13.84 | 14.2 | 14.8 | 15.36 | 14.59 | 14.9 | 15.1994 | 14 |  | 14.25 | 14.45 | 15.75 | 15.1 | 15.5 | 15.47 | 14.5 | 15.1 | 14.3 | 14.1533 | 15.084 | 14.0017 |  |  | 1
| 1 | 2 | 13.9157 | 14.42 | 14.8 | 17.2 | 12.2302 | 15.3 | 15.27 | 14 | 13.54 | 14.1 | 14.4 | 14.65 | 14.27 | 14.78 | 15.3936 | 14 | > 15 | 14.85 | 14.55 | 15.75 | 14.2 | 15.1 | 15.32 | 14.9 | 15 | 14.2 | 13.88 | 15.3767 | 14.0279 |  |  | #N/A
| 1 | 3 | 13.9335 | 14.44 | 14.8 | 17.1 | 12.2553 | 15 | 15.14 | 13.75 | 14.04 | 14.3 | 14.3 | 15.24 | 14.29 | 14.77 | 15.1416 | 14.1 | > 15 | 14.9 | 14.75 | 15.2 | 15 | 15.2 | 14.71 | 14.65 | 15 | 14.2 | 14.02 | 15.0586 | 14.2884 |  |  | 16
| 1 | 4 | 13.8035 | 14.55 | > 15 | 17.3 | 11.9945 | 15.3 | 15.41 | 14.05 | 13.71 | 15.2 | 14.6 | 15.4 | 14.36 | 14.89 | 15.0937 | 13.9 | > 15 | 14.5 | 14.55 | 15.8 | 14.35 | 15.4 | 15.33 | 14.65 | 14.7 | 14.2 | 13.8833 | 15.2232 | 13.9552 |  |  | 14.6498
| 1 | 5 | 14.0673 | 14.56 | > 15 | 17.1 | 12.155 | 15.2 | 15.19 | 14 | 13.91 | 14 | 14.9 | 15.69 | 14.45 | 14.89 | 15.5384 | 13.9 |  | 14.45 | 14.75 | 15.25 | 13.85 | 15.4 | 15.38 | 14.5 | 14.6 | 14.2 | 14.1967 | 14.9465 | 13.9151 |  |  | 97
| 1 | 6 | 13.8395 | 14.45 | 14.6 | 17.4 | 12.3003 | 15.4 | 15.1 | 13.85 | 13.99 | 14.3 | 14.7 | 15.28 | 14.6 | 14.85 | 15.3661 | 13.9 |  | 14.45 | 14.5 | 16.1 | 14.25 | 15.5 | 15.32 | 14.45 | 14.7 | 14 | 14.02 | 14.8608 | 13.9589
| Mean |  | 13.9376 | 14.515 | 14.7333 | 17.2333 | 12.2107 | 15.2333 | 15.1867 | 13.975 | 13.8383 | 14.35 | 14.6167 | 15.27 | 14.4267 | 14.8467 | 15.2888 | 13.9667 |  | 14.5667 | 14.5917 | 15.6417 | 14.4583 | 15.35 | 15.255 | 14.6083 | 14.85 | 14.1833 | 14.0256 | 15.0916 | 14.0245
| Median |  | 13.9246 | 14.5 | 14.8 | 17.25 | 12.2427 | 15.25 | 15.165 | 14 | 13.875 | 14.25 | 14.65 | 15.32 | 14.405 | 14.87 | 15.2828 | 13.95 |  | 14.475 | 14.55 | 15.75 | 14.3 | 15.4 | 15.325 | 14.575 | 14.85 | 14.2 | 14.02 | 15.0713 | 13.9803
| Std Dev. |  | 0.1109 | 0.096073 | 0.11547 | 0.12111 | 0.12179 | 0.13663 | 0.13981 | 0.15732 | 0.1867 | 0.43243 | 0.23166 | 0.34258 | 0.14487 | 0.058195 | 0.17144 | 0.08165 |  | 0.2543 | 0.12813 | 0.34845 | 0.4893 | 0.16432 | 0.27311 | 0.16558 | 0.20736 | 0.098319 | 0.13197 | 0.18658 | 0.1351
| Rel.Std.Dev. |  | 0.7957% | 0.6619% | 0.7837% | 0.7027% | 0.9974% | 0.8969% | 0.9206% | 1.126% | 1.349% | 3.013% | 1.585% | 2.243% | 1.004% | 0.392% | 1.121% | 0.5846% |  | 1.746% | 0.8781% | 2.228% | 3.384% | 1.07% | 1.79% | 1.133% | 1.396% | 0.6932% | 0.9409% | 1.236% | 0.9633%
| PDM3 |  | -4.861% | -0.9201% | 0.5703% | 17.64% | -16.65% | 3.983% | 3.665% | -4.606% | -5.539% | -2.046% | -0.2261% | 4.234% | -1.523% | 1.344% | 4.362% | -4.663% |  | -0.5674% | -0.3967% | 6.771% | -1.307% | 4.78% | 4.131% | -0.283% | 1.367% | -3.184% | -4.261% | 3.016% | -4.268%
| Z-Score (Absolute) |  | 0.78 | 0.1 | 0.16 | 3.09 | 2.8 | 0.74 | 0.69 | 0.73 | 0.89 | 0.29 | 0.02 | 0.79 | 0.2 | 0.29 | 0.81 | 0.74 | NA | 0.04 | 0.01 | 1.22 | 0.17 | 0.88 | 0.77 | 0.01 | 0.29 | 0.49 | 0.67 | 0.58 | 0.67
| Analytical results for Ga in OREAS 45f (Certified Value 26.7 ppm) | [32 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 24.8587 | 27.5 | 26.1 | 25.8 | 26.4814 | 25.2 | 27.19 | 26.7 | 37 | 28.9 | 25.8 | 28 | 25.73 | 25.3211 | 28.6 | 32.7 | 27.6 | 26.5 | 30.1 | 27.9 | 28.1 | 30.3 | 26.1 | 27.7 | 26.42 | 24.208 | 26.1008 |  |  |  |  | 1
| 1 | 2 | 25.003 | 28.8 | 26.2 | 26.3 | 25.1314 | 26 | 28.56 | 26.1 | 37 | 27.3 | 25 | 29 | 26.37 | 25.9122 | 28.5 | 31.1 | 24.1 | 26.9 | 28.5 | 26.7 | 26.1 | 29.8 | 27.3 | 27.9 | 27.97 | 22.437 | 26.4848 |  |  |  |  | #N/A
| 1 | 3 | 24.9267 | 26.6 | 26.9 | 26.4 | 20.5186 | 25.2 | 28.38 | 25.9 | 37 | 28.4 | 25.2 | 26 | 26.45 | 26.7585 | 29 | 31.1 | 24.3 | 26.5 | 28.2 | 26.7 | 26.4 | 29.5 | 27.9 | 26.3 | 26.23 | 22.757 | 27.3725 |  |  |  |  | 16
| 1 | 4 | 25.0882 | 28 | 27.2 | 25 | 23.8292 | 25.8 | 28.04 | 26.1 | 38 | 28.6 | 25.4 | 28 | 26.32 | 26.816 | 29.4 | 32.2 | 24.2 | 26.1 | 28.3 | 26.6 | 26.1 | 30.6 | 28.2 | 25.5 | 27.0433 | 21.263 | 26.8454 |  |  |  |  | 26.7491
| 1 | 5 | 24.8967 | 26.2 | 28 | 25.9 | 21.7499 | 25 | 28.5 | 26.1 | 38 | 26.5 | 26.2 | 28 | 25.35 | 26.1354 | 28.7 | 30 | 27.3 | 25.7 | 28 | 24.6 | 26.5 | 30.7 | 28.2 | 26.4 | 26.5467 | 22.422 | 26.8247 |  |  |  |  | 98
| 1 | 6 | 24.9492 | 26.4 | 27.2 | 26.8 | 21.3806 | 25.2 | 29.49 | 25.9 | 39 | 27.4 | 25 | 29 | 26.19 | 27.344 | 28.5 | 29.1 | 27.1 | 23.6 | 28 | 25.9 | 28 | 30.9 | 26.6 | 25.6 | 25.93 | 22.847 | 26.4936
| Mean |  | 24.9537 | 27.25 | 26.9333 | 26.0333 | 23.1818 | 25.4 | 28.36 | 26.1333 | 37.6667 | 27.85 | 25.4333 | 28 | 26.0683 | 26.3812 | 28.7833 | 31.0333 | 25.7667 | 25.8833 | 28.5167 | 26.4 | 26.8667 | 30.3 | 27.3833 | 26.5667 | 26.69 | 22.6557 | 26.687
| Median |  | 24.938 | 27.05 | 27.05 | 26.1 | 22.7895 | 25.2 | 28.44 | 26.1 | 37.5 | 27.9 | 25.3 | 28 | 26.255 | 26.4469 | 28.65 | 31.1 | 25.7 | 26.3 | 28.25 | 26.65 | 26.45 | 30.45 | 27.6 | 26.35 | 26.4833 | 22.597 | 26.6592
| Std Dev. |  | 0.081942 | 1.0271 | 0.70899 | 0.62183 | 2.345 | 0.4 | 0.74943 | 0.29439 | 0.8165 | 0.92682 | 0.48028 | 1.0954 | 0.43481 | 0.72974 | 0.35449 | 1.3382 | 1.7247 | 1.1907 | 0.79854 | 1.0918 | 0.93095 | 0.54772 | 0.87958 | 1.0231 | 0.72745 | 0.94853 | 0.43264
| Rel.Std.Dev. |  | 0.3284% | 3.769% | 2.632% | 2.389% | 10.12% | 1.575% | 2.643% | 1.127% | 2.168% | 3.328% | 1.888% | 3.912% | 1.668% | 2.766% | 1.232% | 4.312% | 6.694% | 4.6% | 2.8% | 4.136% | 3.465% | 1.808% | 3.212% | 3.851% | 2.726% | 4.187% | 1.621%
| PDM3 |  | -6.712% | 1.873% | 0.6887% | -2.676% | -13.34% | -5.044% | 6.022% | -2.302% | 40.81% | 4.116% | -4.919% | 4.676% | -2.545% | -1.375% | 7.605% | 16.02% | -3.673% | -3.237% | 6.608% | -1.305% | 0.4395% | 13.27% | 2.371% | -0.682% | -0.221% | -15.3% | -0.2323%
| Z-Score (Absolute) |  | 1.27 | 0.41 | 0.18 | 0.48 | 2.57 | 0.94 | 1.23 | 0.41 | 8.04 | 0.85 | 0.92 | 0.96 | 0.45 | 0.22 | 1.54 | 3.18 | 0.67 | 0.59 | 1.34 | 0.21 | 0.13 | 2.65 | 0.51 | 0.09 | 0 | 2.95 | 0
| Analytical results for Gd in OREAS 45f (Certified Value 2.31 ppm) | [32 empty] | Certified
Gd | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab |  |  | 1
| No | No | 02: 02; 04: 12; 05: 13; 07: 15; 10: 18; 12: 23; 13: 24; 14: 27; 17: 33; 18: 38; 19: 39; 20: 40 |  |  | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*OES/MS; 17: 4A*MS; 18: 4A*MS; 19: 5A*OES; 20: 4A*MS |  |  | 2
2
| 1 | 1 | 02: 2.6714; 04: 2.2; 05: 2.35; 07: 1.67; 10: 3.1; 12: 2.31; 13: 2.2377; 14: 2.62; 17: 2.3; 18: 2.27; 19: 1.2; 20: 2.3779 |  |  | 1
| 1 | 2 | 02: 2.6294; 04: 2.2; 05: 2.42; 07: 1.74; 10: 3; 12: 2.33; 13: 2.3655; 14: 2.74; 17: 1.9; 18: 2.38; 19: 1.1; 20: 2.3721 |  |  | #N/A
| 1 | 3 | 02: 2.748; 04: 2; 05: 2.45; 07: 1.97; 10: 3.2; 12: 2.15; 13: 2.3604; 14: 2.63; 17: 2.1; 18: 2.23; 19: 1.1; 20: 2.3245 |  |  | 16
| 1 | 4 | 02: 2.69; 04: 2.2; 05: 2.02; 07: 1.84; 10: 3.2; 12: 2.36; 13: 2.237; 14: 2.62; 17: 2.2; 18: 2.2633; 19: 1.2; 20: 2.2043 |  |  | 2.3082
| 1 | 5 | 02: 2.6723; 04: 2; 05: 2.4; 07: 1.6; 10: 2.8; 12: 2.45; 13: 2.3307; 14: 2.61; 17: 2.5; 18: 2.2633; 19: 1; 20: 2.382 |  |  | 99
| 1 | 6 | 02: 2.656; 04: 2; 05: 2.51; 07: 1.89; 10: 2.7; 12: 2.3; 13: 2.3561; 14: 2.62; 17: 2.5; 18: 2.2667; 19: 1.1; 20: 2.221
| Mean |  | 02: 2.6778; 04: 2.1; 05: 2.3583; 07: 1.785; 10: 3; 12: 2.3167; 13: 2.3145; 14: 2.64; 17: 2.25; 18: 2.2789; 19: 1.1167; 20: 2.3136
| Median |  | 02: 2.6719; 04: 2.1; 05: 2.41; 07: 1.79; 10: 3.05; 12: 2.32; 13: 2.3434; 14: 2.62; 17: 2.25; 18: 2.265; 19: 1.1; 20: 2.3483
| Std Dev. |  | 02: 0.039915; 04: 0.10954; 05: 0.17406; 07: 0.13982; 10: 0.20976; 12: 0.097912; 13: 0.060989; 14: 0.049396; 17: 0.23452; 18: 0.051625; 19: 0.075277; 20: 0.081087
| Rel.Std.Dev. |  | 02: 1.491%; 04: 5.216%; 05: 7.381%; 07: 7.833%; 10: 6.992%; 12: 4.226%; 13: 2.635%; 14: 1.871%; 17: 10.42%; 18: 2.265%; 19: 6.741%; 20: 3.505%
| PDM3 |  | 02: 16.01%; 04: -9.021%; 05: 2.17%; 07: -22.67%; 10: 29.97%; 12: 0.3652%; 13: 0.2735%; 14: 14.37%; 17: -2.523%; 18: -1.271%; 19: -51.62%; 20: 0.2338%
| Z-Score (Absolute) |  | 02: 1.76; 04: 1.04; 05: 0.21; 07: 2.57; 10: 3.33; 12: 0.01; 13: 0; 14: 1.58; 17: 0.31; 18: 0.17; 19: 5.81; 20: 0
| Analytical results for Ge in OREAS 45f (Indicative Value 0.28 ppm) | [32 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 04 | 07 | 10 | 13 | 14 | 21 | 23 | 24 | 27 | 28 | 29 | 30 | 31 | 34 | 35 | 36 | 38 | 39 | [12 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | [12 empty] | 2
2
| 1 | 1 | < 0.25 | 0.5 | 2.6 | 5.2357 | < 0.05 | 0.12 | < 10 | 1.8 | 1.694 | 2.1 | 0.15 | 0.13 | 0.3 | 0.84 | 0.4 | 0.27 | < 0.05 | 0.74 | 0.115 | [12 empty] | 1
| 1 | 2 | < 0.25 | 0.5 | 2.4 | 5.3905 | < 0.05 | 0.1 | < 10 | 1.9 | 1.6808 | 2 | 0.1 | 0.16 | 0.43 | 0.42 | 0.2 | 0.49 | < 0.05 | 0.73667 | 0.115 | [12 empty] | 11
| 1 | 3 | < 0.25 | 0.5 | 2.4 | 5.3951 | 0.07 | 0.11 | < 10 | 2 | 1.7201 | 1.9 | 0.11 | 0.17 | 0.38 | 0.73 | 0.2 | 0.45 | < 0.05 | 0.76333 | 0.115 | [12 empty] | 16
| 1 | 4 | < 0.25 | 0.6 | 2.1 | 5.6291 | 0.08 | 0.11 | < 10 | 2 | 1.7425 | 2 | 0.11 | 0.11 | 0.48 | 0.31 | 0.2 | 0.36 | < 0.05 | 0.54667 | 0.115 | [12 empty] | 0.27889
| 1 | 5 | < 0.25 | 0.6 | 2.2 | 5.7034 | < 0.05 | 0.12 | < 10 | 2.1 | 1.7521 | 2 | 0.13 | 0.18 | 0.34 | 0.7 | 0.2 | 0.35 | < 0.05 | 0.42333 | 0.115 | [12 empty] | 24
| 1 | 6 | < 0.25 | 0.6 | 2.3 | 5.2619 | < 0.05 | 0.12 | < 10 | 1.8 | 1.7375 | 2 | 0.13 | 0.12 | 0.37 | 0.5 | 0.3 | 0.26 | < 0.05 | 0.45 | 0.115
| Mean |  | 04: 0.55; 05: 2.3333; 07: 5.4359; 10: 0.075; 12: 0.11333; 14: 1.9333; 17: 1.7212; 18: 2; 19: 0.12167; 20: 0.145; 21: 0.38333; 23: 0.58333; 24: 0.25; 25: 0.36333; 28: 0.61; 29: 0.115
| Median |  | 04: 0.55; 05: 2.35; 07: 5.3928; 10: 0.075; 12: 0.115; 14: 1.95; 17: 1.7288; 18: 2; 19: 0.12; 20: 0.145; 21: 0.375; 23: 0.6; 24: 0.2; 25: 0.355; 28: 0.64167; 29: 0.115
| Std Dev. |  | 04: 0.054772; 05: 0.17512; 07: 0.19131; 10: 0.0070711; 12: 0.008165; 14: 0.12111; 17: 0.028453; 18: 0.063246; 19: 0.018348; 20: 0.02881; 21: 0.064083; 23: 0.20461; 24: 0.083666; 25: 0.09288; 28: 0.15551; 29: 0
| Rel.Std.Dev. |  | 04: 9.959%; 05: 7.505%; 07: 3.519%; 10: 9.428%; 12: 7.204%; 14: 6.264%; 17: 1.653%; 18: 3.162%; 19: 15.08%; 20: 19.87%; 21: 16.72%; 23: 35.08%; 24: 33.47%; 25: 25.56%; 28: 25.49%; 29: 0%
| PDM3 |  | 04: 97.21%; 05: 736.7%; 07: 1849%; 10: -73.11%; 12: -59.36%; 14: 593.2%; 17: 517.2%; 18: 617.1%; 19: -56.37%; 20: -48.01%; 21: 37.45%; 23: 109.2%; 24: -10.36%; 25: 30.28%; 28: 118.7%; 29: -58.76%
| Z-Score (Absolute) |  | 0.49 | NA | NA | 19.78 | 0.81 | 0.54 | 18.11 | NA | 5.6 | NA | 0.51 | 0.42 | 0.49 | 1.26 | NA | 0.42 | 0.87 | 1.36 | 0.53
Results from laboratories 4, 7, 23, 27 & 34 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Hf in OREAS 45f (Certified Value 4.64 ppm) | [32 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 4.9386 | 5.4 | 4.74 | 5.2 | 5.6028 | 5.2 | 4.59 | 4.13 | 4.98 | 4.95 | 4.4 | 4.5 | 3.8 | 4.26 | 4.4431 | 4.1 | 5.06 | 4.81 | 4.07 | 4.8 | 4.71 | 4.66 | 4.45 | 4.65 | 5 | 4.31 | 4.157 | 6.1665 |  |  |  | 1
| 1 | 2 | 4.8952 | 5.6 | 4.7 | 5.5 | 5.4639 | 5.4 | 5.13 | 4.19 | 4.82 | 4.72 | 4.8 | 4.2 | 4.1 | 4.34 | 4.5455 | 4.21 | 4.93 | 4.7 | 4.21 | 4.7 | 4.68 | 3.94 | 4.64 | 4.76 | 4.8 | 4.54 | 3.859 | 6.2484 |  |  |  | #N/A
| 1 | 3 | 4.8889 | 5.3 | 4.74 | 5.2 | 4.6901 | 5 | 4.69 | 3.94 | 5.41 | 4.8 | 4.4 | 4.4 | 3.8 | 4.21 | 4.6738 | 4.04 | 4.9 | 4.73 | 4.04 | 4.7 | 4.62 | 4.02 | 4.41 | 4.79 | 4.7 | 4.2467 | 4.082 | 6.307 |  |  |  | 16
| 1 | 4 | 4.8914 | 5.8 | 4.84 | 5.2 | 5.1484 | 5.2 | 4.33 | 4.19 | 5.12 | 4.95 | 4.4 | 4.5 | 3.9 | 4.34 | 4.6014 | 4.07 | 4.99 | 4.42 | 4.41 | 4.8 | 4.85 | 4.33 | 4.53 | 5.06 | 4.6 | 4.39 | 3.577 | 6.2373 |  |  |  | 4.6375
| 1 | 5 | 4.9701 | 5.4 | 5.05 | 5.2 | 5.1223 | 4.8 | 4.57 | 4.09 | 4.88 | 4.71 | 4.4 | 4.7 | 3.9 | 4.39 | 4.4875 | 4.21 | 5.48 | 4.4 | 4.13 | 4.9 | 4.35 | 4.56 | 4.62 | 4.95 | 4.8 | 4.4233 | 3.885 | 6.144 |  |  |  | 100
| 1 | 6 | 4.9531 | 5.3 | 4.7 | 5.7 | 5.4696 | 5 | 4.78 | 4.06 | 5.18 | 4.82 | 4.8 | 4.3 | 3.7 | 4.25 | 4.6132 | 4.13 | 4.48 | 4.75 | 4.53 | 4.8 | 4.39 | 4.46 | 4.75 | 4.66 | 4.7 | 4.3333 | 3.891 | 6.3363
| Mean |  | 4.9229 | 5.4667 | 4.795 | 5.3333 | 5.2495 | 5.1 | 4.6817 | 4.1 | 5.065 | 4.825 | 4.5333 | 4.4333 | 3.8667 | 4.2983 | 4.5607 | 4.1267 | 4.9733 | 4.635 | 4.2317 | 4.7833 | 4.6 | 4.3283 | 4.5667 | 4.8117 | 4.7667 | 4.3739 | 3.9085 | 6.2399
| Median |  | 4.9169 | 5.4 | 4.74 | 5.2 | 5.3061 | 5.1 | 4.64 | 4.11 | 5.05 | 4.81 | 4.4 | 4.45 | 3.85 | 4.3 | 4.5735 | 4.115 | 4.96 | 4.715 | 4.17 | 4.8 | 4.65 | 4.395 | 4.575 | 4.775 | 4.75 | 4.3617 | 3.888 | 6.2428
| Std Dev. |  | 0.035502 | 0.19664 | 0.13502 | 0.21602 | 0.33421 | 0.20976 | 0.26656 | 0.094234 | 0.2176 | 0.10597 | 0.20656 | 0.17512 | 0.13663 | 0.068532 | 0.08559 | 0.071181 | 0.32098 | 0.17807 | 0.19722 | 0.075277 | 0.19391 | 0.29219 | 0.12754 | 0.16315 | 0.13663 | 0.10219 | 0.20251 | 0.075444
| Rel.Std.Dev. |  | 0.7212% | 3.597% | 2.816% | 4.05% | 6.367% | 4.113% | 5.694% | 2.298% | 4.296% | 2.196% | 4.556% | 3.95% | 3.533% | 1.594% | 1.877% | 1.725% | 6.454% | 3.842% | 4.661% | 1.574% | 4.215% | 6.751% | 2.793% | 3.391% | 2.866% | 2.336% | 5.181% | 1.209%
| PDM3 |  | 6.154% | 17.88% | 3.396% | 15% | 13.2% | 9.973% | 0.9521% | -11.59% | 9.218% | 4.043% | -2.246% | -4.403% | -16.62% | -7.314% | -1.655% | -11.02% | 7.241% | -0.0542% | -8.751% | 3.144% | -0.8089% | -6.667% | -1.528% | 3.755% | 2.785% | -5.685% | -15.72% | 34.55%
| Z-Score (Absolute) |  | 0.6 | 1.82 | 0.31 | 1.52 | 1.33 | 0.99 | 0.05 | 1.26 | 0.91 | 0.37 | 0.28 | 0.51 | 1.78 | 0.81 | 0.22 | 1.2 | 0.71 | 0.05 | 0.96 | 0.28 | 0.13 | 0.74 | 0.21 | 0.34 | 0.24 | 0.64 | 1.69 | 3.56
| Analytical results for Hg in OREAS 45f (Indicative Value 0.033 ppm) | [32 empty] | Indicative
Hg | Round | Replicate | 02: Lab |  |  | 1
| No | No | 02: 39 |  |  | ppm
|  |  | 02: 5A*MS |  |  | 3
3
| 1 | 1 | 02: 0.026 |  |  | 1
| 1 | 2 | 02: 0.039 |  |  | 19
| 1 | 3 | 02: 0.04 |  |  | 16
| 1 | 4 | 02: 0.021 |  |  | 0.033167
| 1 | 5 | 02: 0.034 |  |  | 25
| 1 | 6 | 02: 0.039
| Mean |  | 02: 0.033167
| Median |  | 02: 0.0365
| Std Dev. |  | 02: 0.0079352
| Rel.Std.Dev. |  | 02: 23.93%
| PDM3 |  | 02: -8.882e-14%
Z-Score (Absolute) | bar: 02
| Analytical results for Ho in OREAS 45f (Certified Value 0.45 ppm) | [32 empty] | Certified
Ho | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab |  |  | 1
| No | No | 02: 02; 04: 10; 05: 12; 07: 13; 10: 18; 12: 23; 13: 24; 14: 27; 17: 33; 18: 38; 19: 39; 20: 40 |  |  | ppm
|  |  | 02: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 07: 4A*OES/MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*OES/MS; 17: 4A*MS; 18: 4A*MS; 19: 5A*MS; 20: 4A*MS |  |  | 2
2
| 1 | 1 | 02: 0.4443; 04: 0.64798; 05: 0.48; 07: 0.48; 10: 0.51; 12: 0.42; 13: 0.42851; 14: 0.4; 17: 0.43; 18: 0.43; 19: 0.369; 20: 0.5408 |  |  | 1
| 1 | 2 | 02: 0.439; 04: 0.61658; 05: 0.48; 07: 0.52; 10: 0.5; 12: 0.44; 13: 0.4509; 14: 0.41; 17: 0.38; 18: 0.44667; 19: 0.331; 20: 0.5584 |  |  | #N/A
| 1 | 3 | 02: 0.4453; 04: 0.53399; 05: 0.46; 07: 0.45; 10: 0.52; 12: 0.43; 13: 0.44368; 14: 0.4; 17: 0.43; 18: 0.42; 19: 0.359; 20: 0.5501 |  |  | 16
| 1 | 4 | 02: 0.4425; 04: 0.61654; 05: 0.46; 07: 0.45; 10: 0.57; 12: 0.43; 13: 0.43351; 14: 0.4; 17: 0.41; 18: 0.43667; 19: 0.316; 20: 0.5237 |  |  | 0.44714
| 1 | 5 | 02: 0.4485; 04: 0.52892; 05: 0.44; 07: 0.52; 10: 0.5; 12: 0.43; 13: 0.43122; 14: 0.42; 17: 0.45; 18: 0.44333; 19: 0.347; 20: 0.5407 |  |  | 101
| 1 | 6 | 02: 0.4254; 04: 0.51248; 05: 0.46; 07: 0.49; 10: 0.49; 12: 0.44; 13: 0.42069; 14: 0.41; 17: 0.47; 18: 0.44; 19: 0.355; 20: 0.5179
| Mean |  | 02: 0.44083; 04: 0.57608; 05: 0.46333; 07: 0.485; 10: 0.515; 12: 0.43167; 13: 0.43475; 14: 0.40667; 17: 0.42833; 18: 0.43611; 19: 0.34617; 20: 0.5386
| Median |  | 02: 0.4434; 04: 0.57526; 05: 0.46; 07: 0.485; 10: 0.505; 12: 0.43; 13: 0.43236; 14: 0.405; 17: 0.43; 18: 0.43833; 19: 0.351; 20: 0.54075
| Std Dev. |  | 02: 0.0081862; 04: 0.057422; 05: 0.015055; 07: 0.031464; 10: 0.02881; 12: 0.0075277; 13: 0.010874; 14: 0.008165; 17: 0.031252; 18: 0.0097563; 19: 0.019518; 20: 0.015389
| Rel.Std.Dev. |  | 02: 1.857%; 04: 9.968%; 05: 3.249%; 07: 6.487%; 10: 5.594%; 12: 1.744%; 13: 2.501%; 14: 2.008%; 17: 7.296%; 18: 2.237%; 19: 5.638%; 20: 2.857%
| PDM3 |  | 02: -1.411%; 04: 28.84%; 05: 3.621%; 07: 8.467%; 10: 15.18%; 12: -3.461%; 13: -2.77%; 14: -9.052%; 17: -4.206%; 18: -2.467%; 19: -22.58%; 20: 20.45%
| Z-Score (Absolute) |  | 02: 0.06; 04: 3.27; 05: 0.59; 07: 1.11; 10: 1.82; 12: 0.16; 13: 0.09; 14: 0.76; 17: 0.24; 18: 0.06; 19: 2.2; 20: 2.38
| Analytical results for I in OREAS 45f (Indicative Value 14.3 ppm) | [32 empty] | Indicative
I | Round | Replicate | 02: Lab |  |  | 1
| No | No | 02: 39 |  |  | ppm
|  |  | 02: 5A*MS |  |  | 1
1
| 1 | 1 | 02: 11 |  |  | 1
| 1 | 2 | 02: 9 |  |  | 12
| 1 | 3 | 02: 14 |  |  | 16
| 1 | 4 | 02: 16 |  |  | 14.3333
| 1 | 5 | 02: 18 |  |  | 21
| 1 | 6 | 02: 18
| Mean |  | 02: 14.3333
| Median |  | 02: 15
| Std Dev. |  | 02: 3.7238
| Rel.Std.Dev. |  | 02: 25.98%
| PDM3 |  | 02: 2.442e-13%
Z-Score (Absolute) | bar: 02
| Analytical results for In in OREAS 45f (Certified Value 0.11 ppm) | [32 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | < 0.25 | 0.11 | 0.11 | 0.09 | 0.12 | 0.099 | 0.104 | < 0.5 | 0.11 | 0.1 | 0.15 | 0.1 | 0.1 | 0.11953 | 0.093 | 0.1 | 0.104 | 0.112 | 0.096 | 0.109 | 0.14 | 0.12 | 0.113 | 0.123 | 0.1 | 0.105 | 0.1054 |  |  |  |  | 1
| 1 | 2 | < 0.25 | 0.11 | 0.11 | 0.13 | 0.12 | 0.118 | 0.11 | < 0.5 | 0.11 | 0.1 | 0.13 | 0.11 | 0.11 | 0.089334 | 0.097 | 0.09 | 0.11 | 0.132 | 0.109 | 0.102 | 0.13 | 0.11 | 0.113 | 0.118 | 0.10333 | 0.093 | 0.1011 |  |  |  |  | 13
| 1 | 3 | < 0.25 | 0.1 | 0.1 | 0.11 | 0.12 | 0.095 | 0.098 | < 0.5 | 0.11 | 0.1 | 0.11 | 0.1 | 0.09 | 0.10888 | 0.094 | 0.1 | 0.115 | 0.1 | 0.112 | 0.094 | 0.12 | 0.11 | 0.099 | 0.117 | 0.1 | 0.105 | 0.109 |  |  |  |  | 16
| 1 | 4 | < 0.25 | 0.12 | 0.11 | 0.1 | 0.12 | 0.127 | 0.104 | < 0.5 | 0.11 | 0.1 | 0.15 | 0.11 | 0.1 | 0.10656 | 0.092 | 0.1 | 0.105 | 0.099 | 0.102 | 0.105 | 0.11 | 0.12 | 0.115 | 0.115 | 0.1 | 0.093 | 0.1014 |  |  |  |  | 0.10554
| 1 | 5 | < 0.25 | 0.1 | 0.11 | 0.09 | 0.12 | 0.094 | 0.101 | < 0.5 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.079332 | 0.099 | 0.1 | 0.108 | 0.106 | 0.11 | 0.079 | 0.12 | 0.11 | 0.106 | 0.114 | 0.1 | 0.097 | 0.1052 |  |  |  |  | 102
| 1 | 6 | < 0.25 | 0.11 | 0.11 | 0.11 | 0.12 | 0.133 | 0.093 | < 0.5 | 0.11 | 0.1 | 0.16 | 0.1 | 0.09 | 0.12263 | 0.092 | 0.09 | 0.106 | 0.11 | 0.112 | 0.097 | 0.14 | 0.12 | 0.125 | 0.109 | 0.1 | 0.097 | 0.1026
| Mean |  |  | 0.10833 | 0.10833 | 0.105 | 0.12 | 0.111 | 0.10167 |  | 0.10833 | 0.1 | 0.13333 | 0.10333 | 0.098333 | 0.10438 | 0.0945 | 0.096667 | 0.108 | 0.10983 | 0.10683 | 0.097667 | 0.12667 | 0.115 | 0.11183 | 0.116 | 0.10056 | 0.098333 | 0.10412
| Median |  |  | 0.11 | 0.11 | 0.105 | 0.12 | 0.1085 | 0.1025 |  | 0.11 | 0.1 | 0.14 | 0.1 | 0.1 | 0.10772 | 0.0935 | 0.1 | 0.107 | 0.108 | 0.1095 | 0.0995 | 0.125 | 0.115 | 0.113 | 0.116 | 0.1 | 0.097 | 0.1039
| Std Dev. |  |  | 0.0075277 | 0.0040825 | 0.015166 | 0 | 0.017193 | 0.0058195 |  | 0.0040825 | 1.5202e-17 | 0.024221 | 0.005164 | 0.0075277 | 0.01698 | 0.002881 | 0.005164 | 0.0040497 | 0.01204 | 0.0064627 | 0.010614 | 0.012111 | 0.0054772 | 0.0087731 | 0.0046476 | 0.0013608 | 0.005465 | 0.0030149
| Rel.Std.Dev. |  |  | 6.949% | 3.768% | 14.44% | 0% | 15.49% | 5.724% |  | 3.768% | 1.52e-14% | 18.17% | 4.997% | 7.655% | 16.27% | 3.049% | 5.342% | 3.75% | 10.96% | 6.049% | 10.87% | 9.561% | 4.763% | 7.845% | 4.007% | 1.353% | 5.558% | 2.896%
| PDM3 |  |  | 2.644% | 2.644% | -0.5147% | 13.7% | 5.17% | -3.673% |  | 2.644% | -5.252% | 26.33% | -2.094% | -6.831% | -1.103% | -10.46% | -8.41% | 2.328% | 4.065% | 1.222% | -7.463% | 20.01% | 8.96% | 5.96% | 9.908% | -4.726% | -6.831% | -1.352%
| Z-Score (Absolute) |  | 1.64 | 0.03 | 0.03 | 0.29 | 1.16 | 0.29 | 0.61 | 13.68 | 0.03 | 0.77 | 2.44 | 0.45 | 0.93 | 0.35 | 1.3 | 1.09 | 0 | 0.18 | 0.11 | 1 | 1.8 | 0.67 | 0.37 | 0.77 | 0.72 | 0.93 | 0.37
| Analytical results for Ir in OREAS 45f (Indicative Value < 300 ppb) | [32 empty] | Indicative
Ir | Round | Replicate | 02: Lab; 04: Lab |  |  | 1
| No | No | 02: 10; 04: 39 |  |  | ppb
|  |  | 02: 4A*OES/MS; 04: 5A*MS |  |  | 0
0
| 1 | 1 | 02: 297.8582; 04: < 5 |  |  | 1
| 1 | 2 | 02: 257.1407; 04: < 5 |  |  | 26
| 1 | 3 | 02: 231.1187; 04: < 5 |  |  | 16
| 1 | 4 | 02: 229.308; 04: < 5 |  |  | < 300
| 1 | 5 | 02: 228.9976; 04: < 5 |  |  | 22
| 1 | 6 | 02: 265.8586; 04: < 5
| Mean |  | 02: 251.7136
| Median |  | 02: 244.1297
| Std Dev. |  | 02: 27.5719
| Rel.Std.Dev. |  | 02: 10.95%
PDM3
| Z-Score (Absolute) |  | 02: 0.67; 04: 0.67
| Analytical results for K in OREAS 45f (Certified Value 0.224 wt.%) | [32 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.23101 | 0.25 | 0.23 | 0.224 | 0.24188 | 0.23 | 0.25 | 0.22 | 0.22 | 0.22 | 0.23 | 0.23 | 0.22 | 0.2231 | 0.21621 | 0.215 | 0.3 | 0.22 | 0.22 | 0.23 | 0.23 | 0.234 | 0.25 | 0.22 | 0.22 | 0.23 | 0.22633 | 0.26324 | 0.21102 |  |  | 1
| 1 | 2 | 0.22446 | 0.24 | 0.22 | 0.224 | 0.24394 | 0.23 | 0.25 | 0.22 | 0.23 | 0.214 | 0.23 | 0.21 | 0.21 | 0.2248 | 0.2097 | 0.216 | 0.3 | 0.22 | 0.22 | 0.24 | 0.22 | 0.217 | 0.23 | 0.23 | 0.21 | 0.23 | 0.217 | 0.27085 | 0.2142 |  |  | #N/A
| 1 | 3 | 0.21308 | 0.24 | 0.23 | 0.221 | 0.23133 | 0.23 | 0.23 | 0.22 | 0.22 | 0.224 | 0.23 | 0.22 | 0.22 | 0.2261 | 0.21076 | 0.221 | 0.29 | 0.22 | 0.23 | 0.23 | 0.23 | 0.224 | 0.23 | 0.22 | 0.21 | 0.23 | 0.21863 | 0.26365 | 0.21687 |  |  | 16
| 1 | 4 | 0.22401 | 0.25 | 0.22 | 0.219 | 0.24026 | 0.23 | 0.26 | 0.22 | 0.23 | 0.226 | 0.23 | 0.23 | 0.21 | 0.2264 | 0.21202 | 0.218 | 0.28 | 0.21 | 0.22 | 0.24 | 0.22 | 0.212 | 0.24 | 0.22 | 0.21 | 0.23 | 0.21877 | 0.26576 | 0.21289 |  |  | 0.22436
| 1 | 5 | 0.22782 | 0.25 | 0.23 | 0.218 | 0.24507 | 0.23 | 0.22 | 0.22 | 0.23 | 0.212 | 0.22 | 0.23 | 0.22 | 0.2251 | 0.21337 | 0.215 | 0.29 | 0.23 | 0.23 | 0.23 | 0.22 | 0.227 | 0.24 | 0.22 | 0.21 | 0.23 | 0.2236 | 0.26185 | 0.21132 |  |  | 103
| 1 | 6 | 0.2342 | 0.24 | 0.23 | 0.223 | 0.24069 | 0.23 | 0.24 | 0.22 | 0.23 | 0.219 | 0.22 | 0.23 | 0.22 | 0.2228 | 0.21471 | 0.215 | 0.28 | 0.22 | 0.22 | 0.24 | 0.22 | 0.225 | 0.24 | 0.22 | 0.21 | 0.23 | 0.22133 | 0.26173 | 0.21305
| Mean |  | 0.22576 | 0.245 | 0.22667 | 0.2215 | 0.24053 | 0.23 | 0.24167 | 0.22 | 0.22667 | 0.21917 | 0.22667 | 0.225 | 0.21667 | 0.22472 | 0.21279 | 0.21667 | 0.29 | 0.22 | 0.22333 | 0.235 | 0.22333 | 0.22317 | 0.23833 | 0.22167 | 0.21167 | 0.23 | 0.22094 | 0.26451 | 0.21323
| Median |  | 0.22614 | 0.245 | 0.23 | 0.222 | 0.24128 | 0.23 | 0.245 | 0.22 | 0.23 | 0.2195 | 0.23 | 0.23 | 0.22 | 0.22495 | 0.2127 | 0.2155 | 0.29 | 0.22 | 0.22 | 0.235 | 0.22 | 0.2245 | 0.24 | 0.22 | 0.21 | 0.23 | 0.22005 | 0.26344 | 0.21297
| Std Dev. |  | 0.007331 | 0.0054772 | 0.005164 | 0.0025884 | 0.0048742 | 0 | 0.01472 | 0 | 0.005164 | 0.0054559 | 0.005164 | 0.0083666 | 0.005164 | 0.0014959 | 0.0024495 | 0.0024221 | 0.0089443 | 0.0063246 | 0.005164 | 0.0054772 | 0.005164 | 0.0077309 | 0.0075277 | 0.0040825 | 0.0040825 | 0 | 0.0035181 | 0.0034301 | 0.0021379
| Rel.Std.Dev. |  | 3.247% | 2.236% | 2.278% | 1.169% | 2.026% | 0% | 6.091% | 0% | 2.278% | 2.489% | 2.278% | 3.718% | 2.383% | 0.6657% | 1.151% | 1.118% | 3.084% | 2.875% | 2.312% | 2.331% | 2.312% | 3.464% | 3.158% | 1.842% | 1.929% | 0% | 1.592% | 1.297% | 1.003%
| PDM3 |  | 0.6258% | 9.2% | 1.029% | -1.274% | 7.207% | 2.515% | 7.715% | -1.943% | 1.029% | -2.314% | 1.029% | 0.2861% | -3.428% | 0.1598% | -5.154% | -3.428% | 29.26% | -1.943% | -0.4568% | 4.743% | -0.4568% | -0.5311% | 6.229% | -1.2% | -5.657% | 2.515% | -1.522% | 17.9% | -4.961%
| Z-Score (Absolute) |  | 0.13 | 2.59 | 0.25 | 0.41 | 2.02 | 0.67 | 2.16 | 0.6 | 0.25 | 0.71 | 0.25 | 0.04 | 1.03 | 0 | 1.52 | 1.03 | 8.33 | 0.6 | 0.18 | 1.31 | 0.18 | 0.2 | 1.74 | 0.39 | 1.67 | 0.67 | 0.48 | 5.08 | 1.47
| Analytical results for La in OREAS 45f (Certified Value 15.7 ppm) | [32 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.418 | 15 | 16.5 | 29.8401 | 16.8 | 16.4 | 11.95 | 18 | 16 | 22.3 | 18.2 | 16.1 | 15.62 | 16.7693 | 14.6 | 14.2 | 16.05 | 16.4 | 20.2 | 15.55 | 17.2 | 14.2 | 15.5 | 18 | 15.4 | 14.8233 | 15.138 | 14.8139 |  |  |  | 1
| 1 | 2 | 14.413 | 15.1 | 16.4 | 26.9631 | 17.4 | 17 | 14.3 | 18 | 9 | 23.6 | 15.8 | 17.1 | 15.83 | 16.6216 | 14.4 | 15.7 | 14.75 | 16.5 | 16.6 | 15.6 | 18.8 | 15.7 | 15.65 | 17.9 | 14.7 | 15.6867 | 14.97 | 14.9239 |  |  |  | #N/A
| 1 | 3 | 13.954 | 14 | 16 | 24.6014 | 17 | 17 | 12.15 | 18 | 13 | 21.7 | 16.2 | 15.2 | 15.99 | 16.243 | 14.2 | 14.3 | 14.65 | 16.2 | 16.8 | 15.65 | 16.1 | 13.2 | 15.9 | 17.1 | 14.6 | 14.5667 | 14.664 | 14.8647 |  |  |  | 16
| 1 | 4 | 14.772 | 15.3 | 16.4 | 30.7502 | 17.1 | 15.7 | 13.95 | 18 | 11 | 17.5 | 17 | 17.6 | 15.89 | 16.3793 | 14.6 | 13.6 | 13.15 | 14.9 | 16.8 | 15.55 | 18.2 | 12.8 | 15.8 | 16.6 | 15.1 | 15.1167 | 14.847 | 13.6215 |  |  |  | 15.7192
| 1 | 5 | 15.646 | 15.1 | 16.7 | 25.7629 | 16.6 | 16.8 | 10.35 | 18 | 12 | 16.9 | 18.6 | 16.1 | 16.05 | 17.004 | 14.7 | 13.5 | 13.9 | 16.75 | 16.9 | 14.3 | 16.4 | 15.1 | 15.9 | 17.3 | 14.9 | 14.67 | 14.746 | 15.0946 |  |  |  | 104
| 1 | 6 | 16.049 | 14.9 | 15.1 | 30.3927 | 16.8 | 17.2 | 8.72 | 18 | 11 | 15.8 | 17.7 | 16.2 | 16.08 | 16.8653 | 14.4 | 12.7 | 14.25 | 15 | 16.8 | 14.3 | 17.1 | 15.2 | 15.4 | 17.1 | 15.1 | 14.8267 | 14.417 | 14.7833
| Mean |  | 14.8753 | 14.9 | 16.1833 | 28.0517 | 16.95 | 16.6833 | 11.9033 | 18 | 12 | 19.6333 | 17.25 | 16.3833 | 15.91 | 16.6471 | 14.4833 | 14 | 14.4583 | 15.9583 | 17.35 | 15.1583 | 17.3 | 14.3667 | 15.6917 | 17.3333 | 14.9667 | 14.9483 | 14.797 | 14.6837
| Median |  | 14.595 | 15.05 | 16.4 | 28.4016 | 16.9 | 16.9 | 12.05 | 18 | 11.5 | 19.6 | 17.35 | 16.15 | 15.94 | 16.6955 | 14.5 | 13.9 | 14.45 | 16.3 | 16.8 | 15.55 | 17.15 | 14.65 | 15.725 | 17.2 | 15 | 14.825 | 14.7965 | 14.8393
| Std Dev. |  | 0.8067 | 0.46043 | 0.57764 | 2.6187 | 0.28107 | 0.55287 | 2.1233 | 0 | 2.3664 | 3.2813 | 1.1131 | 0.84715 | 0.17076 | 0.29175 | 0.18348 | 1.0119 | 0.97181 | 0.80151 | 1.3996 | 0.6659 | 1.0354 | 1.1708 | 0.21075 | 0.53166 | 0.29439 | 0.4067 | 0.25036 | 0.53187
| Rel.Std.Dev. |  | 5.423% | 3.09% | 3.569% | 9.335% | 1.658% | 3.314% | 17.84% | 0% | 19.72% | 16.71% | 6.453% | 5.171% | 1.073% | 1.753% | 1.267% | 7.228% | 6.721% | 5.023% | 8.067% | 4.393% | 5.985% | 8.149% | 1.343% | 3.067% | 1.967% | 2.721% | 1.692% | 3.622%
| PDM3 |  | -5.368% | -5.212% | 2.953% | 78.46% | 7.83% | 6.133% | -24.28% | 14.51% | -23.66% | 24.9% | 9.738% | 4.225% | 1.214% | 5.903% | -7.862% | -10.94% | -8.021% | 1.521% | 10.37% | -3.568% | 10.06% | -8.604% | -0.1752% | 10.27% | -4.787% | -4.904% | -5.867% | -6.588%
| Z-Score (Absolute) |  | 0.55 | 0.54 | 0.23 | 7.29 | 0.68 | 0.53 | 2.32 | 1.31 | 2.26 | 2.28 | 0.86 | 0.35 | 0.06 | 0.5 | 0.78 | 1.07 | 0.8 | 0.09 | 0.92 | 0.38 | 0.89 | 0.85 | 0.06 | 0.91 | 0.5 | 0.51 | 0.6 | 0.66
| Analytical results for Li in OREAS 45f (Certified Value 20.4 ppm) | [32 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  | 1
2
| 1 | 1 | 21.339 | 21 | 21 | 20.2 | 45.7225 | 19 | 19 | 17.6 | 18 | 21 | 22.8 | 20 | 20.7 | 19.7 | 19.8 | 20.3322 | 22.9 | 24 | 18.1 | 20.4 | 20.7 | 20.7 | 20.8 | 22 | 21 | 21.1 | 21 | 19.2 | 22.4599 | 18.0388 |  | 1
| 1 | 2 | 21.174 | 22 | 20 | 21.5 | 45.9299 | 20 | 20.1 | 17.9 | 18 | 21 | 21.1 | 20 | 20.4 | 20.1 | 20.3 | 20.7218 | 22.4 | 23 | 28.1 | 21.1 | 20.3 | 19.8 | 19.6 | 21 | 24.2 | 21.2 | 20 | 19.9333 | 20.0586 | 18.751 |  | 21
| 1 | 3 | 20.981 | 18 | 20 | 20.6 | 48.0756 | 18 | 20.3 | 14.7 | 18 | 21 | 22.3 | 20 | 20 | 19.5 | 19.8 | 21.0909 | 22.6 | 24 | 34.3 | 20.7 | 19.8 | 19 | 21 | 20 | 28.4 | 20.8 | 21 | 19.4333 | 21.7299 | 19.5074 |  | 16
| 1 | 4 | 21.309 | 22 | 20 | 20.3 | 45.6783 | 20 | 20.7 | 17.2 | 18 | 21 | 21.6 | 20 | 21.8 | 20.3 | 20.2 | 20.606 | 22.5 | 24 | 30.1 | 20.5 | 21.4 | 20.1 | 18.5 | 22 | 34.9 | 20.4 | 20 | 19.3667 | 18.81 | 19.1191 |  | 20.4363
| 1 | 5 | 21.068 | 21 | 21 | 20.3 | 48.6273 | 18.5 | 19.6 | 17.7 | 18 | 21 | 21.4 | 20 | 22.3 | 19 | 19.9 | 20.7518 | 23.1 | 24 | 18.6 | 20.1 | 19.4 | 17.6 | 21 | 22 | 33.8 | 20.5 | 20 | 19.6 | 21.5649 | 19.0635 |  | 105
| 1 | 6 | 21.338 | 21 | 21 | 20 | 51.3964 | 18.5 | 21.6 | 17 | 18 | 21 | 22.1 | 20 | 21 | 20.2 | 19.9 | 21.1312 | 22.7 | 23 | 20.4 | 18.8 | 20.3 | 18.3 | 21.2 | 22 | 25.6 | 20.6 | 21 | 18.7 | 20.4824 | 18.7027
| Mean |  | 21.2015 | 20.8333 | 20.5 | 20.4833 | 47.5717 | 19 | 20.2167 | 17.0167 | 18 | 21 | 21.8833 | 20 | 21.0333 | 19.8 | 19.9833 | 20.7723 | 22.7 | 23.6667 | 24.9333 | 20.2667 | 20.3167 | 19.25 | 20.35 | 21.5 | 27.9833 | 20.7667 | 20.5 | 19.3722 | 20.8509 | 18.8637
| Median |  | 21.2415 | 21 | 20.5 | 20.3 | 47.0028 | 18.75 | 20.2 | 17.4 | 18 | 21 | 21.85 | 20 | 20.85 | 19.9 | 19.9 | 20.7368 | 22.65 | 24 | 24.25 | 20.45 | 20.3 | 19.4 | 20.9 | 22 | 27 | 20.7 | 20.5 | 19.4 | 21.0236 | 18.9073
| Std Dev. |  | 0.15251 | 1.472 | 0.54772 | 0.53448 | 2.267 | 0.83666 | 0.89759 | 1.1822 | 0 | 0 | 0.63061 | 0 | 0.86872 | 0.498 | 0.2137 | 0.30153 | 0.26077 | 0.5164 | 6.809 | 0.79162 | 0.6969 | 1.1675 | 1.0728 | 0.83666 | 5.4883 | 0.3266 | 0.54772 | 0.4128 | 1.3267 | 0.49771
| Rel.Std.Dev. |  | 0.7193% | 7.065% | 2.672% | 2.609% | 4.765% | 4.403% | 4.44% | 6.947% | 0% | 0% | 2.882% | 0% | 4.13% | 2.515% | 1.069% | 1.452% | 1.149% | 2.182% | 27.31% | 3.906% | 3.43% | 6.065% | 5.272% | 3.891% | 19.61% | 1.573% | 2.672% | 2.131% | 6.363% | 2.638%
| PDM3 |  | 3.744% | 1.943% | 0.3118% | 0.2302% | 132.8% | -7.028% | -1.075% | -16.73% | -11.92% | 2.758% | 7.081% | -2.135% | 2.922% | -3.113% | -2.216% | 1.644% | 11.08% | 15.81% | 22.01% | -0.83% | -0.5853% | -5.805% | -0.4222% | 5.205% | 36.93% | 1.617% | 0.3118% | -5.207% | 2.029% | -7.695%
| Z-Score (Absolute) |  | 0.68 | 0.32 | 0 | 0.02 | 26.08 | 1.44 | 0.27 | 3.36 | 2.41 | 0.48 | 1.33 | NA | 0.51 | 0.67 | 0.5 | 0.26 | 2.12 | 3.05 | 4.27 | 0.22 | 0.18 | 1.2 | 0.14 | 0.96 | 7.21 | 0.26 | 0 | 1.09 | 0.34 | 1.58
Results from laboratory 19 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 45f (Certified Value 0.19 ppm) | [32 empty] | Certified
Lu | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab |  | 1
| No | No | 02: 02; 04: 04; 05: 05; 07: 12; 10: 13; 12: 15; 13: 18; 14: 23; 15: 24; 17: 25; 18: 27; 19: 33; 20: 34; 21: 38; 23: 39; 24: 40 |  | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*OES/MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*MS; 15: 4A*OES/MS; 17: 4A*MS; 18: 4A*OES/MS; 19: 4A*MS; 20: 4A*MS; 21: 4A*MS; 23: 5A*MS; 24: 4A*MS |  | 2
2
| 1 | 1 | 02: < 0.25; 04: 0.16; 05: 0.19; 07: 0.16; 10: 0.19; 12: 0.2; 13: 0.2; 14: 0.16; 15: 0.1956; 17: 0.148; 18: 0.17; 19: 0.18; 20: 0.15; 21: 0.17667; 23: 0.159; 24: 0.2477 |  | 1
| 1 | 2 | 02: < 0.25; 04: 0.18; 05: 0.19; 07: 0.18; 10: 0.21; 12: 0.18; 13: 0.19; 14: 0.17; 15: 0.20193; 17: 0.172; 18: 0.18; 19: 0.18; 20: 0.16; 21: 0.18667; 23: 0.142; 24: 0.2582 |  | #N/A
| 1 | 3 | 02: < 0.25; 04: 0.16; 05: 0.21; 07: 0.16; 10: 0.2; 12: 0.2; 13: 0.21; 14: 0.2; 15: 0.20399; 17: 0.152; 18: 0.19; 19: 0.2; 20: 0.15; 21: 0.18333; 23: 0.149; 24: 0.2537 |  | 16
| 1 | 4 | 02: < 0.25; 04: 0.17; 05: 0.2; 07: 0.18; 10: 0.19; 12: 0.22; 13: 0.24; 14: 0.18; 15: 0.19381; 17: 0.151; 18: 0.19; 19: 0.21; 20: 0.15; 21: 0.18333; 23: 0.131; 24: 0.2447 |  | 0.18593
| 1 | 5 | 02: < 0.25; 04: 0.17; 05: 0.21; 07: 0.16; 10: 0.21; 12: 0.2; 13: 0.21; 14: 0.17; 15: 0.19827; 17: 0.167; 18: 0.18; 19: 0.23; 20: 0.15; 21: 0.18333; 23: 0.147; 24: 0.2491 |  | 106
| 1 | 6 | 02: < 0.25; 04: 0.18; 05: 0.19; 07: 0.16; 10: 0.21; 12: 0.2; 13: 0.2; 14: 0.16; 15: 0.20037; 17: 0.155; 18: 0.18; 19: 0.22; 20: 0.15; 21: 0.18333; 23: 0.145; 24: 0.2344
| Mean |  | 04: 0.17; 05: 0.19833; 07: 0.16667; 10: 0.20167; 12: 0.2; 13: 0.20833; 14: 0.17333; 15: 0.199; 17: 0.1575; 18: 0.18167; 19: 0.20333; 20: 0.15167; 21: 0.18278; 23: 0.1455; 24: 0.24797
| Median |  | 04: 0.17; 05: 0.195; 07: 0.16; 10: 0.205; 12: 0.2; 13: 0.205; 14: 0.17; 15: 0.19932; 17: 0.1535; 18: 0.18; 19: 0.205; 20: 0.15; 21: 0.18333; 23: 0.146; 24: 0.2484
| Std Dev. |  | 04: 0.0089443; 05: 0.0098319; 07: 0.010328; 10: 0.0098319; 12: 0.012649; 13: 0.017224; 14: 0.015055; 15: 0.0038555; 17: 0.0096902; 18: 0.0075277; 19: 0.020656; 20: 0.0040825; 21: 0.0032773; 23: 0.0091597; 24: 0.0081691
| Rel.Std.Dev. |  | 04: 5.261%; 05: 4.957%; 07: 6.197%; 10: 4.875%; 12: 6.325%; 13: 8.268%; 14: 8.686%; 15: 1.938%; 17: 6.153%; 18: 4.144%; 19: 10.16%; 20: 2.692%; 21: 1.793%; 23: 6.295%; 24: 3.294%
| PDM3 |  | 04: -8.568%; 05: 6.671%; 07: -10.36%; 10: 8.463%; 12: 7.567%; 13: 12.05%; 14: -6.775%; 15: 7.026%; 17: -15.29%; 18: -2.293%; 19: 9.36%; 20: -18.43%; 21: -1.696%; 23: -21.75%; 24: 33.37%
| Z-Score (Absolute) |  | 02: 2.07; 04: 0.44; 05: 0.58; 07: 0.56; 10: 0.7; 12: 0.64; 13: 0.95; 14: 0.32; 15: 0.61; 17: 0.9; 18: 0.02; 19: 0.76; 20: 1.11; 21: 0.02; 23: 1.33; 24: 2.38
| Analytical results for Mg in OREAS 45f (Certified Value 0.229 wt.%) | [32 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.2444 | 0.24 | 0.24 | 0.269 | 0.2386 | 0.24 | 0.26 | 0.21 | 0.21 | 0.23 | 0.23 | 0.24 | 0.22 | 0.2251 | 0.22409 | 0.225 | 0.26 | 0.21 | 0.21 | 0.23 | 0.22 | 0.247 | 0.24 | 0.22 | 0.25 | 0.23 | 0.22663 | 0.31071 | 0.20175 |  |  | 1
| 1 | 2 | 0.24363 | 0.22 | 0.24 | 0.259 | 0.2349 | 0.24 | 0.24 | 0.21 | 0.13 | 0.232 | 0.23 | 0.22 | 0.2 | 0.2235 | 0.23027 | 0.224 | 0.25 | 0.22 | 0.21 | 0.23 | 0.21 | 0.226 | 0.24 | 0.22 | 0.25 | 0.22 | 0.21897 | 0.31681 | 0.19861 |  |  | #N/A
| 1 | 3 | 0.24247 | 0.23 | 0.24 | 0.253 | 0.23515 | 0.23 | 0.24 | 0.2 | 0.19 | 0.231 | 0.24 | 0.22 | 0.21 | 0.2235 | 0.22534 | 0.23 | 0.26 | 0.22 | 0.22 | 0.23 | 0.22 | 0.241 | 0.22 | 0.22 | 0.24 | 0.22 | 0.21957 | 0.31128 | 0.20765 |  |  | 16
| 1 | 4 | 0.24118 | 0.24 | 0.24 | 0.25 | 0.23115 | 0.25 | 0.25 | 0.21 | 0.17 | 0.232 | 0.23 | 0.22 | 0.19 | 0.2246 | 0.229 | 0.224 | 0.26 | 0.21 | 0.21 | 0.24 | 0.22 | 0.233 | 0.23 | 0.22 | 0.25 | 0.22 | 0.22173 | 0.32156 | 0.18859 |  |  | 0.22932
| 1 | 5 | 0.2464 | 0.24 | 0.25 | 0.252 | 0.2329 | 0.24 | 0.23 | 0.21 | 0.19 | 0.219 | 0.23 | 0.24 | 0.21 | 0.2247 | 0.23134 | 0.223 | 0.25 | 0.21 | 0.22 | 0.23 | 0.21 | 0.241 | 0.24 | 0.22 | 0.24 | 0.22 | 0.22483 | 0.31781 | 0.20914 |  |  | 107
| 1 | 6 | 0.2481 | 0.23 | 0.25 | 0.251 | 0.2361 | 0.24 | 0.23 | 0.2 | 0.17 | 0.228 | 0.24 | 0.22 | 0.21 | 0.2233 | 0.22916 | 0.228 | 0.25 | 0.21 | 0.22 | 0.24 | 0.21 | 0.257 | 0.24 | 0.22 | 0.25 | 0.22 | 0.2244 | 0.31415 | 0.19865
| Mean |  | 0.24436 | 0.23333 | 0.24333 | 0.25567 | 0.2348 | 0.24 | 0.24167 | 0.20667 | 0.17667 | 0.22867 | 0.23333 | 0.22667 | 0.20667 | 0.22412 | 0.2282 | 0.22567 | 0.255 | 0.21333 | 0.215 | 0.23333 | 0.215 | 0.24083 | 0.235 | 0.22 | 0.24667 | 0.22167 | 0.22269 | 0.31539 | 0.20073
| Median |  | 0.24402 | 0.235 | 0.24 | 0.2525 | 0.23503 | 0.24 | 0.24 | 0.21 | 0.18 | 0.2305 | 0.23 | 0.22 | 0.21 | 0.22405 | 0.22908 | 0.2245 | 0.255 | 0.21 | 0.215 | 0.23 | 0.215 | 0.241 | 0.24 | 0.22 | 0.25 | 0.22 | 0.22307 | 0.31548 | 0.2002
| Std Dev. |  | 0.0025443 | 0.008165 | 0.005164 | 0.0072572 | 0.0025766 | 0.0063246 | 0.01169 | 0.005164 | 0.027325 | 0.0049666 | 0.005164 | 0.010328 | 0.010328 | 0.0007705 | 0.002853 | 0.0027325 | 0.0054772 | 0.005164 | 0.0054772 | 0.005164 | 0.0054772 | 0.010778 | 0.0083666 | 0 | 0.005164 | 0.0040825 | 0.0030855 | 0.0041523 | 0.0074284
| Rel.Std.Dev. |  | 1.041% | 3.499% | 2.122% | 2.839% | 1.097% | 2.635% | 4.837% | 2.499% | 15.47% | 2.172% | 2.213% | 4.556% | 4.997% | 0.3438% | 1.25% | 1.211% | 2.148% | 2.421% | 2.548% | 2.213% | 2.548% | 4.475% | 3.56% | 0% | 2.094% | 1.842% | 1.386% | 1.317% | 3.701%
| PDM3 |  | 6.56% | 1.75% | 6.11% | 11.49% | 2.389% | 4.657% | 5.384% | -9.879% | -22.96% | -0.2853% | 1.75% | -1.157% | -9.879% | -2.269% | -0.4888% | -1.594% | 11.2% | -6.972% | -6.245% | 1.75% | -6.245% | 5.02% | 2.476% | -4.065% | 7.564% | -3.338% | -2.892% | 37.53% | -12.47%
| Z-Score (Absolute) |  | 0.87 | 0.26 | 0.81 | 1.5 | 0.34 | 0.63 | 0.72 | 1.22 | 2.88 | 0 | 0.26 | 0.11 | 1.22 | 0.25 | 0.03 | 0.17 | 1.46 | 0.85 | 0.76 | 0.26 | 0.76 | 0.67 | 0.35 | 0.48 | 1 | 0.39 | 0.33 | 4.81 | 1.55
| Analytical results for Mn in OREAS 45f (Certified Value 0.022 wt.%) | [32 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.023047 | 0.0261 | 0.0228 | 0.022 | 0.0202 | 0.0244 | 0.025 | 0.0216 | 0.0191 | 0.035 | 0.0222 | 0.0242 | 0.0243 | 0.0214 | 0.0219 | 0.022185 | 0.0223 | 0.0208 | 0.0219 | 0.0234 | 0.0227 | 0.023 | 0.03 | 0.0234 | 0.0206 | 0.0217 | 0.0216 | 0.01978 | 0.021493 |  |  | 1
| 1 | 2 | 0.023162 | 0.0269 | 0.0219 | 0.0222 | 0.02045 | 0.0242 | 0.0233 | 0.0212 | 0.0195 | 0.035 | 0.0214 | 0.0248 | 0.0233 | 0.0217 | 0.0216 | 0.023445 | 0.0218 | 0.021 | 0.0212 | 0.0246 | 0.0225 | 0.0229 | 0.03 | 0.0242 | 0.0204 | 0.0216 | 0.020833 | 0.019484 | 0.021679 |  |  | #N/A
| 1 | 3 | 0.023124 | 0.0254 | 0.0219 | 0.0218 | 0.02005 | 0.0238 | 0.0226 | 0.021 | 0.0197 | 0.036 | 0.0226 | 0.0232 | 0.0241 | 0.0218 | 0.0217 | 0.022499 | 0.0232 | 0.0213 | 0.0221 | 0.0232 | 0.0231 | 0.0229 | 0.03 | 0.0235 | 0.0201 | 0.0207 | 0.0213 | 0.019929 | 0.021888 |  |  | 16
| 1 | 4 | 0.02322 | 0.027 | 0.0221 | 0.0214 | 0.01995 | 0.024 | 0.0235 | 0.0214 | 0.0195 | 0.035 | 0.0227 | 0.0242 | 0.0245 | 0.0225 | 0.0216 | 0.02269 | 0.0221 | 0.0207 | 0.0211 | 0.0237 | 0.0226 | 0.0214 | 0.03 | 0.0237 | 0.0206 | 0.0216 | 0.0213 | 0.020028 | 0.021495 |  |  | 0.022116
| 1 | 5 | 0.022745 | 0.0257 | 0.0222 | 0.0218 | 0.02025 | 0.0242 | 0.0225 | 0.0212 | 0.0191 | 0.035 | 0.0215 | 0.0236 | 0.0242 | 0.0216 | 0.022 | 0.022789 | 0.0223 | 0.0216 | 0.0229 | 0.0237 | 0.0211 | 0.0229 | 0.03 | 0.0234 | 0.0204 | 0.0216 | 0.0215 | 0.019652 | 0.021543 |  |  | 108
| 1 | 6 | 0.023012 | 0.0248 | 0.0219 | 0.0222 | 0.0203 | 0.0248 | 0.0221 | 0.0208 | 0.0197 | 0.036 | 0.0218 | 0.0246 | 0.024 | 0.0226 | 0.0223 | 0.023341 | 0.022 | 0.0211 | 0.0223 | 0.024 | 0.0221 | 0.0241 | 0.03 | 0.0235 | 0.0206 | 0.021 | 0.0216 | 0.01945 | 0.021567
| Mean |  | 0.023051 | 0.025983 | 0.022133 | 0.0219 | 0.0202 | 0.024233 | 0.023167 | 0.0212 | 0.019433 | 0.035333 | 0.022033 | 0.0241 | 0.024067 | 0.021933 | 0.02185 | 0.022825 | 0.022283 | 0.021083 | 0.021917 | 0.023767 | 0.02235 | 0.022867 | 0.03 | 0.023617 | 0.02045 | 0.021367 | 0.021356 | 0.019721 | 0.021611
| Median |  | 0.023085 | 0.0259 | 0.022 | 0.0219 | 0.020225 | 0.0242 | 0.02295 | 0.0212 | 0.0195 | 0.035 | 0.022 | 0.0242 | 0.02415 | 0.02175 | 0.0218 | 0.022739 | 0.0222 | 0.02105 | 0.022 | 0.0237 | 0.02255 | 0.0229 | 0.03 | 0.0235 | 0.0205 | 0.0216 | 0.0214 | 0.019716 | 0.021555
| Std Dev. |  | 0.00016823 | 0.0008612 | 0.00035024 | 0.00030332 | 0.00017889 | 0.00034448 | 0.0010386 | 0.00028284 | 0.00027325 | 0.0005164 | 0.00055377 | 0.00060332 | 0.00041312 | 0.00049666 | 0.00027386 | 0.00048714 | 0.00048751 | 0.00033116 | 0.0006824 | 0.00049261 | 0.0006921 | 0.00085946 | 0 | 0.00030605 | 0.00019748 | 0.00041312 | 0.00028957 | 0.00023464 | 0.00015176
| Rel.Std.Dev. |  | 0.7298% | 3.314% | 1.582% | 1.385% | 0.8856% | 1.422% | 4.483% | 1.334% | 1.406% | 1.462% | 2.513% | 2.503% | 1.717% | 2.264% | 1.253% | 2.134% | 2.188% | 1.571% | 3.114% | 2.073% | 3.097% | 3.759% | 0% | 1.296% | 0.9657% | 1.933% | 1.356% | 1.19% | 0.7022%
| PDM3 |  | 4.228% | 17.48% | 0.0765% | -0.9785% | -8.665% | 9.572% | 4.749% | -4.144% | -12.13% | 59.76% | -0.3757% | 8.969% | 8.818% | -0.8278% | -1.205% | 3.203% | 0.7547% | -4.671% | -0.9032% | 7.462% | 1.056% | 3.392% | 35.65% | 6.783% | -7.535% | -3.39% | -3.44% | -10.83% | -2.285%
| Z-Score (Absolute) |  | 0.71 | 2.84 | 0.04 | 0.13 | 1.37 | 1.57 | 0.79 | 0.64 | 1.93 | 9.65 | 0.04 | 1.47 | 1.44 | 0.11 | 0.17 | 0.54 | 0.15 | 0.73 | 0.12 | 1.23 | 0.19 | 0.57 | NA | 1.12 | 1.19 | 0.52 | 0.53 | 1.72 | 0.34
Results from laboratory 34 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Mo in OREAS 45f (Certified Value 2.27 ppm) | [32 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.5816 | 2.33 | 2.3 | 2.1 | 3.3278 | 2.5 | 2.29 | 2.17 | 1.1 | < 2 | 2.1 | 2.5 | 2.4 | 2.2 | 2.1 | 2.1508 | 2.32 | 2.55 | 2.19 | 2.12 | 2.51 | 2.23 | 2.3 | 2.52 | 2.28 | 2.55 | 2 | 2.1333 | 1.824 | 2.4388 |  | 1
| 1 | 2 | 2.5469 | 2.44 | 2.33 | 2.2 | 3.9583 | 2.5 | 2.37 | 2.08 | 1 | < 2 | 1.8 | 2.5 | 2.5 | 2.4 | 2 | 2.2173 | 2.25 | 2.65 | 2.06 | 2.31 | 2.44 | 2.17 | 2.2 | 2.51 | 2.27 | 2.54 | 2 | 2.3 | 1.798 | 2.4806 |  | 8
| 1 | 3 | 2.6084 | 2.35 | 2.32 | 2.1 | 3.4685 | 2.5 | 2.37 | 2.05 | 0.6 | < 2 | 2 | 2.5 | 2.5 | 2.1 | 1.9 | 2.1215 | 2.33 | 2.37 | 1.94 | 2.18 | 2.28 | 2.18 | 2.2 | 2.36 | 2.29 | 2.47 | 2 | 2.1 | 1.806 | 2.518 |  | 16
| 1 | 4 | 2.6987 | 2.29 | 2.47 | 2 | 3.4449 | 2.5 | 2.09 | 2.06 | 1 | < 2 | 1.8 | 3.5 | 2.3 | 2.3 | 2 | 2.1969 | 2.16 | 2.38 | 2.04 | 2.11 | 3.17 | 2.13 | 1.9 | 2.44 | 2.19 | 2.5 | 2 | 2.1667 | 2.222 | 2.4938 |  | 2.2702
| 1 | 5 | 2.6077 | 2.21 | 2.43 | 1.9 | 3.6677 | 2.5 | 2.31 | 2.08 | 0.7 | < 2 | 1.9 | 3 | 2.4 | 2 | 1.8 | 2.1277 | 2.29 | 2.35 | 2.22 | 2.23 | 2.23 | 2.15 | 2.2 | 2.4 | 2.2 | 2.55 | 2 | 2.2 | 1.882 | 2.5522 |  | 109
| 1 | 6 | 2.6934 | 2.36 | 2.4 | 2.2 | 3.9234 | 2.5 | 2.58 | 2.05 | 1.2 | < 2 | 1.8 | 2.5 | 2 | 2.1 | 1.9 | 2.1549 | 2.29 | 2.48 | 2.25 | 2.32 | 2.64 | 2.13 | 2.4 | 2.58 | 2.26 | 2.5 | 2 | 2.1 | 1.829 | 2.5009
| Mean |  | 2.6228 | 2.33 | 2.375 | 2.0833 | 3.6318 | 2.5 | 2.335 | 2.0817 | 0.93333 |  | 1.9 | 2.75 | 2.35 | 2.1833 | 1.95 | 2.1615 | 2.2733 | 2.4633 | 2.1167 | 2.2117 | 2.545 | 2.165 | 2.2 | 2.4683 | 2.2483 | 2.5183 | 2 | 2.1667 | 1.8935 | 2.4974
| Median |  | 2.608 | 2.34 | 2.365 | 2.1 | 3.5681 | 2.5 | 2.34 | 2.07 | 1 |  | 1.85 | 2.5 | 2.4 | 2.15 | 1.95 | 2.1528 | 2.29 | 2.43 | 2.125 | 2.205 | 2.475 | 2.16 | 2.2 | 2.475 | 2.265 | 2.52 | 2 | 2.15 | 1.8265 | 2.4973
| Std Dev. |  | 0.061055 | 0.076681 | 0.068337 | 0.1169 | 0.26343 | 0 | 0.15821 | 0.04535 | 0.23381 |  | 0.12649 | 0.41833 | 0.18708 | 0.1472 | 0.10488 | 0.038123 | 0.062183 | 0.11928 | 0.12176 | 0.091086 | 0.34098 | 0.037815 | 0.16733 | 0.082563 | 0.042622 | 0.033116 | 0 | 0.076012 | 0.16359 | 0.037865
| Rel.Std.Dev. |  | 2.328% | 3.291% | 2.877% | 5.611% | 7.253% | 0% | 6.776% | 2.179% | 25.05% |  | 6.657% | 15.21% | 7.961% | 6.742% | 5.379% | 1.764% | 2.735% | 4.842% | 5.753% | 4.118% | 13.4% | 1.747% | 7.606% | 3.345% | 1.896% | 1.315% | 0% | 3.508% | 8.64% | 1.516%
| PDM3 |  | 15.53% | 2.635% | 4.618% | -8.23% | 59.98% | 10.12% | 2.856% | -8.304% | -58.89% |  | -16.31% | 21.14% | 3.516% | -3.825% | -14.1% | -4.787% | 0.1392% | 8.509% | -6.762% | -2.577% | 12.11% | -4.633% | -3.091% | 8.729% | -0.962% | 10.93% | -11.9% | -4.559% | -16.59% | 10.01%
| Z-Score (Absolute) |  | 1.51 | 0.33 | 0.51 | 0.67 | 5.6 | 1.02 | 0.35 | 0.67 | 5.32 | 5.05 | 1.41 | 2.03 | 0.41 | 0.26 | 1.21 | 0.35 | 0.1 | 0.87 | 0.53 | 0.15 | 1.2 | 0.34 | 0.2 | 0.89 | 0 | 1.09 | NA | 0.33 | 1.44 | 1.01
Results from laboratoriy 37 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 45f (Certified Value 0.063 wt.%) | [32 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [5 empty] | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 5A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.072582 | 0.08 | 0.063 | 0.06685 | 0.07 | 0.07 | 0.062 | 0.07 | 0.076 | 0.07 | 0.06 | 0.057 | 0.0646 | 0.038073 | 0.0625 | 0.014 | 0.06 | 0.04 | 0.069 | 0.059 | 0.09 | 0.061 | 0.07 | 0.067 | 0.076083 | 0.064134 | [5 empty] | 1
| 1 | 2 | 0.072353 | 0.07 | 0.062 | 0.0655 | 0.07 | 0.08 | 0.061 | 0.1 | 0.057 | 0.08 | 0.056 | 0.059 | 0.0643 | 0.036816 | 0.0621 | 0.012 | 0.062 | 0.04 | 0.064 | 0.0545 | 0.11 | 0.064 | 0.07 | 0.0647 | 0.076347 | 0.063526 | [5 empty] | #N/A
| 1 | 3 | 0.072379 | 0.07 | 0.059 | 0.0664 | 0.07 | 0.07 | 0.061 | 0.08 | 0.063 | 0.07 | 0.056 | 0.058 | 0.0636 | 0.04057 | 0.0653 | 0.013 | 0.062 | 0.04 | 0.067 | 0.057 | 0.08 | 0.063 | 0.07 | 0.064833 | 0.077073 | 0.066135 | [5 empty] | 16
| 1 | 4 | 0.072528 | 0.07 | 0.058 | 0.0646 | 0.07 | 0.09 | 0.061 | 0.07 | 0.0595 | 0.08 | 0.054 | 0.061 | 0.0633 | 0.042324 | 0.0627 | 0.012 | 0.063 | 0.04 | 0.065 | 0.054 | 0.09 | 0.062 | 0.07 | 0.065 | 0.076996 | 0.063739 | [5 empty] | 0.062532
| 1 | 5 | 0.072908 | 0.07 | 0.059 | 0.0653 | 0.07 | 0.07 | 0.061 | 0.09 | 0.0585 | 0.07 | 0.059 | 0.058 | 0.0637 | 0.03908 | 0.0621 | 0.015 | 0.062 | 0.04 | 0.063 | 0.057 | 0.11 | 0.062 | 0.07 | 0.0672 | 0.074807 | 0.063526 | [5 empty] | 110
| 1 | 6 | 0.073 | 0.07 | 0.059 | 0.06635 | 0.07 | 0.07 | 0.061 | 0.08 | 0.061 | 0.07 | 0.06 | 0.061 | 0.0644 | 0.04044 | 0.062 | 0.014 | 0.063 | 0.04 | 0.064 | 0.0585 | 0.08 | 0.061 | 0.07 | 0.0661 | 0.075578 | 0.064209
| Mean |  | 0.072625 | 0.071667 | 0.06 | 0.065833 | 0.07 | 0.075 | 0.061167 | 0.081667 | 0.0625 | 0.073333 | 0.0575 | 0.059 | 0.063983 | 0.03955 | 0.062783 | 0.013333 | 0.062 | 0.04 | 0.065333 | 0.056667 | 0.093333 | 0.062167 | 0.07 | 0.065806 | 0.076147 | 0.064212
| Median |  | 0.072555 | 0.07 | 0.059 | 0.065925 | 0.07 | 0.07 | 0.061 | 0.08 | 0.06025 | 0.07 | 0.0575 | 0.0585 | 0.064 | 0.03976 | 0.0623 | 0.0135 | 0.062 | 0.04 | 0.0645 | 0.057 | 0.09 | 0.062 | 0.07 | 0.06555 | 0.076215 | 0.063936
| Std Dev. |  | 0.00027073 | 0.0040825 | 0.002 | 0.00084123 | 0 | 0.0083666 | 0.00040825 | 0.01169 | 0.0069282 | 0.005164 | 0.00251 | 0.0016733 | 0.00051929 | 0.00197 | 0.0012624 | 0.0012111 | 0.0010954 | 0 | 0.0022509 | 0.0020412 | 0.013663 | 0.001169 | 0 | 0.0011202 | 0.00086507 | 0.00098696
| Rel.Std.Dev. |  | 0.3728% | 5.696% | 3.333% | 1.278% | 0% | 11.16% | 0.6674% | 14.31% | 11.09% | 7.042% | 4.365% | 2.836% | 0.8116% | 4.981% | 2.011% | 9.083% | 1.767% | 0% | 3.445% | 3.602% | 14.64% | 1.881% | 0% | 1.702% | 1.136% | 1.537%
| PDM3 |  | 16.14% | 14.61% | -4.05% | 5.279% | 11.94% | 19.94% | -2.184% | 30.6% | -0.05195% | 17.27% | -8.048% | -5.649% | 2.32% | -36.75% | 0.4011% | -78.68% | -0.8515% | -36.03% | 4.479% | -9.38% | 49.26% | -0.585% | 11.94% | 5.234% | 21.77% | 2.685%
| Z-Score (Absolute) |  | 2.19 | NA | 0.5 | 0.75 | NA | NA | 0.25 | NA | 0.04 | NA | 1.03 | 0.71 | 0.35 | 4.85 | 0.1 | 10.43 | 0.07 | NA | 0.64 | 1.21 | NA | 0.04 | NA | 0.74 | 2.94 | 0.4
Results from laboratories 5, 12, 13, 17, 19, 30, 34 & 36 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Nb in OREAS 45f (Certified Value 23.1 ppm) | [32 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 25.3032 | 22.8 | 22.3 | 20.6 | 25.6219 | 18 | 23.06 | 21.7 | 25.8 |  | 23.4 | 18.5 | 22.3 | 25.7 | 22.52 | 22.5382 | 22.4 | 21.3 | 22.8 | 21.8 | 24.7 | 22.2 | 23.5 | 24.6 | 23 | 28.8 | 22.75 | 20.017 | 23.6563 |  |  | 1
| 1 | 2 | 25.1703 | 23.1 | 21.9 | 22 | 28.5468 | 18 | 25.41 | 21.7 | 26.8 |  | 22.2 | 19 | 21.6 | 27.3 | 22.67 | 23.525 | 22.1 | 21.8 | 22.6 | 22.7 | 23.8 | 22.1 | 22.7 | 24.5 | 22.9 | 25.7 | 23.9633 | 18.429 | 24.3187 |  |  | 22
| 1 | 3 | 25.4933 | 21.1 | 22.4 | 21.6 | 24.978 | 16 | 23.33 | 20.9 | 24.7 |  | 22.4 | 16 | 21.5 | 25.3 | 21.72 | 23.3181 | 21.4 | 21.9 | 22.6 | 21.8 | 23.9 | 21.9 | 22.2 | 23.9 | 22.9 | 26.2 | 22.56 | 19.08 | 24.292 |  |  | 16
| 1 | 4 | 24.7481 | 22.7 | 22.8 | 21.1 | 26.6739 | 18 | 23.89 | 21.6 | 27.4 |  | 22.4 | 18.5 | 22.8 | 26.6 | 22.95 | 22.9316 | 21.7 | 22.5 | 21.4 | 21.3 | 24.8 | 21.7 | 22.2 | 25.3 | 22.3 | 27.9 | 22.92 | 17.465 | 24.1518 |  |  | 23.1488
| 1 | 5 | 25.1402 | 22.6 | 23.4 | 20.5 | 26.764 | 17 | 23.51 | 21.3 | 20.1 |  | 22.2 | 18 | 22.8 | 24.8 | 22.71 | 22.984 | 22.5 | 21.6 | 21.5 | 22.3 | 23.6 | 21.2 | 22.6 | 25.1 | 22.6 | 27.3 | 23.0733 | 18.41 | 24.0074 |  |  | 111
| 1 | 6 | 24.8022 | 22.1 | 23 | 21.5 | 25.0106 | 19 | 24.17 | 20.8 | 27.3 |  | 21.7 | 16 | 22.1 | 23.6 | 22.28 | 23.7051 | 22.1 | 21.8 | 22.2 | 21.9 | 23.8 | 20.6 | 23.8 | 25.6 | 22.7 | 26.3 | 22.47 | 18.847 | 24.0894
| Mean |  | 25.1096 | 22.4 | 22.6333 | 21.2167 | 26.2659 | 17.6667 | 23.895 | 21.3333 | 25.35 |  | 22.3833 | 17.6667 | 22.1833 | 25.55 | 22.475 | 23.167 | 22.0333 | 21.8167 | 22.1833 | 21.9667 | 24.1 | 21.6167 | 22.8333 | 24.8333 | 22.7333 | 27.0333 | 22.9561 | 18.708 | 24.0859
| Median |  | 25.1553 | 22.65 | 22.6 | 21.3 | 26.1479 | 18 | 23.7 | 21.45 | 26.3 |  | 22.3 | 18.25 | 22.2 | 25.5 | 22.595 | 23.151 | 22.1 | 21.8 | 22.4 | 21.85 | 23.85 | 21.8 | 22.65 | 24.85 | 22.8 | 26.8 | 22.835 | 18.638 | 24.1206
| Std Dev. |  | 0.28793 | 0.71554 | 0.53914 | 0.59133 | 1.3613 | 1.0328 | 0.8409 | 0.40332 | 2.7675 |  | 0.56006 | 1.3292 | 0.56362 | 1.3126 | 0.43103 | 0.43005 | 0.41793 | 0.39707 | 0.60139 | 0.48028 | 0.51381 | 0.61128 | 0.67132 | 0.6186 | 0.2582 | 1.1793 | 0.54134 | 0.84651 | 0.24156
| Rel.Std.Dev. |  | 1.147% | 3.194% | 2.382% | 2.787% | 5.183% | 5.846% | 3.519% | 1.891% | 10.92% |  | 2.502% | 7.524% | 2.541% | 5.137% | 1.918% | 1.856% | 1.897% | 1.82% | 2.711% | 2.186% | 2.132% | 2.828% | 2.94% | 2.491% | 1.136% | 4.362% | 2.358% | 4.525% | 1.003%
| PDM3 |  | 8.47% | -3.235% | -2.227% | -8.346% | 13.47% | -23.68% | 3.224% | -7.842% | 9.509% |  | -3.307% | -23.68% | -4.171% | 10.37% | -2.911% | 0.07878% | -4.819% | -5.754% | -4.171% | -5.107% | 4.109% | -6.618% | -1.363% | 7.277% | -1.795% | 16.78% | -0.8322% | -19.18% | 4.048%
| Z-Score (Absolute) |  | 1.6 | 0.1 | 0.05 | 0.84 | 2.32 | 3.05 | 0.84 | 0.76 | 1.75 | NA | 0.11 | 3.05 | 0.23 | 1.87 | 0.05 | 0.38 | 0.33 | 0.46 | 0.23 | 0.37 | 0.97 | 0.59 | 0.17 | 1.42 | 0.11 | 2.8 | 0.25 | 2.4 | 0.96
| Analytical results for Nd in OREAS 45f (Certified Value 12.3 ppm) | [32 empty] | Certified
Nd | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab |  | 1
| No | No | 02: 02; 04: 10; 05: 12; 07: 13; 10: 15; 12: 18; 13: 23; 14: 24; 15: 27; 17: 33; 18: 38; 19: 39; 20: 40 |  | ppm
|  |  | 02: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 07: 4A*OES/MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*OES/MS; 17: 4A*MS; 18: 4A*MS; 19: 5A*OES; 20: 4A*MS |  | 1
2
| 1 | 1 | 02: 14.022; 04: 19.9259; 05: 11.8; 07: 12.09; 10: 11.5; 12: 18.7; 13: 12.43; 14: 13.1022; 15: 11.4; 17: 12.8; 18: 12.1133; 19: 15.577; 20: 11.8637 |  | 1
| 1 | 2 | 02: 13.946; 04: 17.6497; 05: 12.4; 07: 12.97; 10: 11; 12: 18.2; 13: 12.07; 14: 12.984; 15: 12.1; 17: 10.3; 18: 12.8333; 19: 15.4; 20: 12.0344 |  | #N/A
| 1 | 3 | 02: 14.064; 04: 16.4197; 05: 11.9; 07: 12.97; 10: 13.3; 12: 18.9; 13: 12.43; 14: 12.7982; 15: 11.2; 17: 12; 18: 12.0533; 19: 15.218; 20: 11.3977 |  | 16
| 1 | 4 | 02: 14.013; 04: 18.5544; 05: 12.1; 07: 11.66; 10: 12.4; 12: 17.7; 13: 12.11; 14: 13.0152; 15: 11.2; 17: 10.7; 18: 12.45; 19: 15.328; 20: 10.7572 |  | 12.2541
| 1 | 5 | 02: 13.993; 04: 17.4063; 05: 12; 07: 12.29; 10: 11; 12: 16.2; 13: 12.42; 14: 12.8245; 15: 10.9; 17: 13.1; 18: 12.1967; 19: 15.423; 20: 12.2852 |  | 112
| 1 | 6 | 02: 13.955; 04: 18.0226; 05: 11.8; 07: 12.92; 10: 11.5; 12: 15.8; 13: 12.54; 14: 12.8117; 15: 10.2; 17: 13.2; 18: 12.1633; 19: 15.032; 20: 11.5063
| Mean |  | 02: 13.9988; 04: 17.9964; 05: 12; 07: 12.4833; 10: 11.7833; 12: 17.5833; 13: 12.3333; 14: 12.9226; 15: 11.1667; 17: 12.0167; 18: 12.3017; 19: 15.3297; 20: 11.6408
| Median |  | 02: 14.003; 04: 17.8362; 05: 11.95; 07: 12.605; 10: 11.5; 12: 17.95; 13: 12.425; 14: 12.9043; 15: 11.2; 17: 12.4; 18: 12.18; 19: 15.364; 20: 11.685
| Std Dev. |  | 02: 0.044115; 04: 1.1823; 05: 0.22804; 07: 0.55396; 10: 0.90203; 12: 1.3014; 13: 0.19398; 14: 0.12805; 15: 0.62183; 17: 1.2545; 18: 0.29393; 19: 0.18757; 20: 0.54365
| Rel.Std.Dev. |  | 02: 0.3151%; 04: 6.57%; 05: 1.9%; 07: 4.438%; 10: 7.655%; 12: 7.401%; 13: 1.573%; 14: 0.9909%; 15: 5.569%; 17: 10.44%; 18: 2.389%; 19: 1.224%; 20: 4.67%
| PDM3 |  | 02: 14.24%; 04: 46.86%; 05: -2.073%; 07: 1.871%; 10: -3.842%; 12: 43.49%; 13: 0.6467%; 14: 5.456%; 15: -8.874%; 17: -1.937%; 18: 0.3883%; 19: 25.1%; 20: -5.005%
| Z-Score (Absolute) |  | 02: 1.91; 04: 6.48; 05: 0.38; 07: 0.17; 10: 0.63; 12: 6.01; 13: 0; 14: 0.67; 15: 1.33; 17: 0.36; 18: 0.04; 19: 3.43; 20: 0.79
| Analytical results for Ni in OREAS 45f (Certified Value 256 ppm) | [32 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*MS | 0
0
| 1 | 1 | 254.7811 | 252.4 | 260 | 274 | 225.5701 | 270 | 259.6 | 253 | 221 | 240 | 264 | 295 | 275.5 | 274.1 | 265 | 269.9882 | 230 | 252.2898 | 252 | 240 | 256 | 277 | 266 | 276 | 220.1 | 255 | 248 | 248 | 255.9667 | 218.407 | 263.2133 | 1
| 1 | 2 | 254.0579 | 252.2 | 249 | 273 | 219.2368 | 266 | 278.3 | 252 | 215 | 250 | 253 | 285 | 264 | 276.2 | 261 | 277.0705 | 229 | 257.4819 | 247 | 245 | 256 | 275 | 251 | 223 | 227.2 | 266 | 247 | 245 | 246.2333 | 217.379 | 261.2916 | #N/A
| 1 | 3 | 255.4074 | 239.1 | 258 | 270 | 221.0936 | 270 | 266.3 | 247 | 222 | 250 | 263 | 290 | 270.5 | 271.6 | 261 | 273.533 | 233 | 253.0319 | 250 | 247 | 259 | 269 | 264 | 258 | 215.8 | 261 | 252 | 241 | 251.4667 | 221.265 | 264.7601 | 16
| 1 | 4 | 254.6229 | 249.5 | 259 | 269 | 226.6467 | 266 | 266.6 | 251 | 217 | 250 | 272 | 275 | 272.7 | 288.4 | 262 | 272.7848 | 231 | 253.5797 | 253 | 239 | 252 | 276 | 253 | 249 | 225.2 | 261 | 248 | 252 | 252.3667 | 222.393 | 264.4926 | 256.4411
| 1 | 5 | 255.1708 | 251.1 | 263 | 273 | 224.8671 | 260 | 266 | 250 | 222 | 250 | 258 | 270 | 277.5 | 274.7 | 264 | 276.6008 | 228 | 261.6213 | 245 | 248 | 266 | 268 | 246 | 253 | 223 | 255 | 253 | 247 | 251.3667 | 220.887 | 267.4206 | 113
| 1 | 6 | 257.2558 | 254 | 251 | 276 | 220.4697 | 272 | 263.8 | 246 | 217 | 250 | 258 | 280 | 274.5 | 289 | 268 | 280.6336 | 232 | 261.0305 | 246 | 248 | 255 | 279 | 251 | 270 | 220.1 | 256 | 248 | 241 | 253.5667 | 217.685 | 265.2782
| Mean |  | 255.216 | 249.7167 | 256.6667 | 272.5 | 222.9807 | 267.3333 | 266.7667 | 249.8333 | 219 | 248.3333 | 261.3333 | 282.5 | 272.45 | 279 | 263.5 | 275.1018 | 230.5 | 256.5059 | 248.8333 | 244.5 | 257.3333 | 274 | 255.1667 | 254.8333 | 221.9 | 259 | 249.3333 | 245.6667 | 251.8278 | 219.6693 | 264.4094
| Median |  | 254.976 | 251.65 | 258.5 | 273 | 222.9804 | 268 | 266.15 | 250.5 | 219 | 250 | 260.5 | 282.5 | 273.6 | 275.45 | 263 | 275.0669 | 230.5 | 255.5308 | 248.5 | 246 | 256 | 275.5 | 252 | 255.5 | 221.55 | 258.5 | 248 | 246 | 251.9167 | 219.647 | 264.6264
| Std Dev. |  | 1.1029 | 5.4109 | 5.465 | 2.5884 | 3.085 | 4.3205 | 6.2292 | 2.7869 | 3.0332 | 4.0825 | 6.5625 | 9.3541 | 4.7807 | 7.6611 | 2.7386 | 3.7604 | 1.8708 | 4.1473 | 3.3116 | 4.0373 | 4.8028 | 4.4721 | 7.9854 | 18.6485 | 4.0959 | 4.4272 | 2.5033 | 4.274 | 3.2275 | 2.1083 | 2.0544
| Rel.Std.Dev. |  | 0.4321% | 2.167% | 2.129% | 0.9499% | 1.384% | 1.616% | 2.335% | 1.115% | 1.385% | 1.644% | 2.511% | 3.311% | 1.755% | 2.746% | 1.039% | 1.367% | 0.8116% | 1.617% | 1.331% | 1.651% | 1.866% | 1.632% | 3.129% | 7.318% | 1.846% | 1.709% | 1.004% | 1.74% | 1.282% | 0.9597% | 0.777%
| PDM3 |  | -0.4777% | -2.622% | 0.08797% | 6.262% | -13.05% | 4.247% | 4.026% | -2.577% | -14.6% | -3.162% | 1.908% | 10.16% | 6.243% | 8.797% | 2.753% | 7.277% | -10.12% | 0.02526% | -2.967% | -4.656% | 0.3479% | 6.847% | -0.497% | -0.6269% | -13.47% | 0.9979% | -2.772% | -4.202% | -1.799% | -14.34% | 3.107%
| Z-Score (Absolute) |  | 0 | 0.4 | 0.11 | 1.27 | 2.36 | 0.89 | 0.85 | 0.39 | 2.66 | 0.5 | 0.45 | 2 | 1.26 | 1.74 | 0.61 | 1.46 | 1.81 | 0.09 | 0.47 | 0.79 | 0.16 | 1.38 | 0 | 0.03 | 2.44 | 0.28 | 0.43 | 0.7 | 0.25 | 2.61 | 0.67
| Analytical results for P in OREAS 45f (Certified Value 0.03 wt.%) | [32 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 04 | 05 | 07 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.03 | 0.03 | 0.029 | 0.035 | 0.0283 | 0.029 | 0.0291 | 0.035 | 0.0165 | 0.025 | 0.031 | 0.028 | 0.0308 | 0.027605 | 0.0297 | 0.0359 | 0.028 | 0.029 | 0.03 | 0.03 | 0.0295 | 0.0323 | 0.034 | 0.0273 | 0.02 | 0.028867 | 0.0335 | 0.027796 |  |  |  | 1
| 1 | 2 | 0.03 | 0.03 | 0.029 | 0.035 | 0.0297 | 0.028 | 0.0298 | 0.034 | 0.0155 | 0.025 | 0.029 | 0.028 | 0.0307 | 0.025651 | 0.0297 | 0.0335 | 0.031 | 0.029 | 0.031 | 0.031 | 0.0275 | 0.0321 | 0.034 | 0.0266 | 0.02 | 0.0279 | 0.034114 | 0.028835 |  |  |  | #N/A
| 1 | 3 | 0.03 | 0.03 | 0.028 | 0.035 | 0.0283 | 0.028 | 0.0294 | 0.033 | 0.0165 | 0.025 | 0.03 | 0.029 | 0.0303 | 0.028728 | 0.0284 | 0.0336 | 0.029 | 0.03 | 0.029 | 0.032 | 0.03 | 0.0338 | 0.032 | 0.0342 | 0.02 | 0.028667 | 0.0353 | 0.029024 |  |  |  | 16
| 1 | 4 | 0.02 | 0.03 | 0.029 | 0.035 | 0.0295 | 0.027 | 0.0297 | 0.034 | 0.0155 | 0.03 | 0.031 | 0.029 | 0.0306 | 0.028012 | 0.0293 | 0.0358 | 0.028 | 0.028 | 0.031 | 0.03 | 0.0305 | 0.0347 | 0.033 | 0.0271 | 0.02 | 0.0288 | 0.036349 | 0.02867 |  |  |  | 0.029632
| 1 | 5 | 0.02 | 0.03 | 0.028 | 0.035 | 0.0285 | 0.028 | 0.0282 | 0.034 | 0.0155 | 0.025 | 0.031 | 0.028 | 0.0306 | 0.029263 | 0.0292 | 0.0319 | 0.029 | 0.03 | 0.03 | 0.028 | 0.0295 | 0.0315 | 0.032 | 0.034 | 0.02 | 0.028 | 0.036891 | 0.028255 |  |  |  | 114
| 1 | 6 | 0.02 | 0.04 | 0.029 | 0.03 | 0.028 | 0.028 | 0.0295 | 0.034 | 0.0165 | 0.025 | 0.03 | 0.029 | 0.0307 | 0.025734 | 0.0296 | 0.0336 | 0.029 | 0.029 | 0.032 | 0.03 | 0.0315 | 0.0341 | 0.032 | 0.0341 | 0.02 | 0.027733 | 0.03339 | 0.028279
| Mean |  | 0.025 | 0.031667 | 0.028667 | 0.034167 | 0.028717 | 0.028 | 0.029283 | 0.034 | 0.016 | 0.025833 | 0.030333 | 0.0285 | 0.030617 | 0.027499 | 0.029317 | 0.03405 | 0.029 | 0.029167 | 0.0305 | 0.030167 | 0.02975 | 0.033083 | 0.032833 | 0.03055 | 0.02 | 0.028328 | 0.034924 | 0.028476
| Median |  | 0.025 | 0.03 | 0.029 | 0.035 | 0.0284 | 0.028 | 0.02945 | 0.034 | 0.016 | 0.025 | 0.0305 | 0.0285 | 0.03065 | 0.027809 | 0.02945 | 0.0336 | 0.029 | 0.029 | 0.0305 | 0.03 | 0.02975 | 0.03305 | 0.0325 | 0.03065 | 0.02 | 0.028333 | 0.034707 | 0.028474
| Std Dev. |  | 0.0054772 | 0.0040825 | 0.0005164 | 0.0020412 | 0.00070545 | 0.00063246 | 0.00058452 | 0.00063246 | 0.00054772 | 0.0020412 | 0.0008165 | 0.00054772 | 0.00017224 | 0.0015118 | 0.00049565 | 0.0015372 | 0.0010954 | 0.00075277 | 0.0010488 | 0.0013292 | 0.0013323 | 0.0012844 | 0.00098319 | 0.003896 | 0 | 0.00050439 | 0.0014883 | 0.00045099
| Rel.Std.Dev. |  | 21.91% | 12.89% | 1.801% | 5.974% | 2.457% | 2.259% | 1.996% | 1.86% | 3.423% | 7.902% | 2.692% | 1.922% | 0.5626% | 5.498% | 1.691% | 4.515% | 3.777% | 2.581% | 3.439% | 4.406% | 4.478% | 3.882% | 2.994% | 12.75% | 0% | 1.781% | 4.261% | 1.584%
| PDM3 |  | -15.63% | 6.867% | -3.257% | 15.3% | -3.088% | -5.507% | -1.176% | 14.74% | -46% | -12.82% | 2.368% | -3.819% | 3.324% | -7.197% | -1.063% | 14.91% | -2.132% | -1.57% | 2.93% | 1.805% | 0.3991% | 11.65% | 10.8% | 3.099% | -32.5% | -4.401% | 17.86% | -3.899%
| Z-Score (Absolute) |  | 2.37 | NA | 0.35 | 2.68 | 0.32 | 0.72 | 0.01 | 2.59 | 7.32 | 1.91 | 0.57 | 0.44 | 0.72 | 0.99 | 0.01 | 2.61 | 0.17 | 0.07 | 0.66 | 0.48 | 0.25 | 2.08 | 1.94 | 0.69 | NA | 0.54 | 3.1 | 0.45
Results from laboratories 5 & 37 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 45f (Certified Value 14.7 ppm) | [32 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  | 1
2
| 1 | 1 | 22.573 | 16.4 | 15.1 | 14 | 26.882 | 14 | 14.8 | 13.25 | 15 | 11 | 15 | 14 | 16.1 | 15 | 13.2 | 13.8536 | 15.3 | 16.2 | 14.25 | 13.7 | 16.9 | 15.05 | 17 | 14.3 | 14.3 | 15.2 | 13 | 13.6 | 9.301 | 14.5966 |  | 1
| 1 | 2 | 22.475 | 17.6 | 14.8 | 14 | 29.9123 | 15 | 15.5 | 13.1 | 13 | 12 | 14 | 13 | 16 | 16 | 13.3 | 14.1862 | 14.8 | 16 | 14.1 | 14.1 | 16.6 | 14.85 | 15 | 14.5 | 14.6 | 15.1 | 14 | 16.9333 | 8.7 | 14.9228 |  | 9
| 1 | 3 | 22.537 | 16.8 | 14.6 | 13 | 28.4481 | 14 | 16.7 | 12.7 | 13 | 13 | 14 | 13 | 15.8 | 15 | 14.4 | 14.3643 | 15.1 | 16.3 | 14 | 13.65 | 16.2 | 14.5 | 18 | 14 | 14.65 | 15.3 | 15 | 13.7333 | 9.869 | 14.4595 |  | 16
| 1 | 4 | 22.392 | 17.4 | 15.4 | 13 | 32.3872 | 15 | 14 | 13.15 | 13 | 12 | 15 | 13 | 16.2 | 15 | 14.6 | 14.4934 | 15 | 16.3 | 13.7 | 14.3 | 21.8 | 14.55 | 14 | 14.4 | 15.1 | 14.3 | 14 | 14.1333 | 9.144 | 14.6831 |  | 14.7159
| 1 | 5 | 22.418 | 17.7 | 15.7 | 13 | 33.324 | 14 | 14.6 | 12.75 | 14 | 11 | 13 | 14 | 16.6 | 16 | 13.3 | 13.8601 | 15.3 | 16.3 | 14 | 13.8 | 17.9 | 14.45 | 16 | 14.3 | 15.2 | 14.9 | 15 | 13.8333 | 9.334 | 14.1562 |  | 115
| 1 | 6 | 22.549 | 16.5 | 15 | 14 | 31.4405 | 15 | 15.8 | 12.9 | 14 | 11 | 14 | 14 | 15.8 | 15 | 12.9 | 14.1958 | 15.2 | 16.2 | 14.35 | 14.45 | 18.1 | 14.25 | 16 | 15.2 | 14.45 | 14.9 | 14 | 14.1 | 8.715 | 14.4851
| Mean |  | 22.4907 | 17.0667 | 15.1 | 13.5 | 30.399 | 14.5 | 15.2333 | 12.975 | 13.6667 | 11.6667 | 14.1667 | 13.5 | 16.0833 | 15.3333 | 13.6167 | 14.1589 | 15.1167 | 16.2167 | 14.0667 | 14 | 17.9167 | 14.6083 | 16 | 14.45 | 14.7167 | 14.95 | 14.1667 | 14.3889 | 9.1772 | 14.5505
| Median |  | 22.506 | 17.1 | 15.05 | 13.5 | 30.6764 | 14.5 | 15.15 | 13 | 13.5 | 11.5 | 14 | 13.5 | 16.05 | 15 | 13.3 | 14.191 | 15.15 | 16.25 | 14.05 | 13.95 | 17.4 | 14.525 | 16 | 14.35 | 14.625 | 15 | 14 | 13.9667 | 9.2225 | 14.5409
| Std Dev. |  | 0.074282 | 0.57155 | 0.4 | 0.54772 | 2.4497 | 0.54772 | 0.96471 | 0.22528 | 0.8165 | 0.8165 | 0.75277 | 0.54772 | 0.29944 | 0.5164 | 0.70261 | 0.26027 | 0.19408 | 0.1169 | 0.2273 | 0.33317 | 2.041 | 0.29055 | 1.4142 | 0.40373 | 0.3587 | 0.35637 | 0.75277 | 1.2636 | 0.43839 | 0.25548
| Rel.Std.Dev. |  | 0.3303% | 3.349% | 2.649% | 4.057% | 8.059% | 3.777% | 6.333% | 1.736% | 5.974% | 6.999% | 5.314% | 4.057% | 1.862% | 3.368% | 5.16% | 1.838% | 1.284% | 0.7209% | 1.616% | 2.38% | 11.39% | 1.989% | 8.839% | 2.794% | 2.437% | 2.384% | 5.314% | 8.782% | 4.777% | 1.756%
| PDM3 |  | 52.83% | 15.97% | 2.61% | -8.262% | 106.6% | -1.467% | 3.516% | -11.83% | -7.13% | -20.72% | -3.732% | -8.262% | 9.292% | 4.196% | -7.47% | -3.785% | 2.724% | 10.2% | -4.412% | -4.865% | 21.75% | -0.7308% | 8.726% | -1.807% | 0.005375% | 1.591% | -3.732% | -2.222% | -37.64% | -1.123%
| Z-Score (Absolute) |  | 7.09 | 2.26 | 0.51 | 0.91 | 14.12 | 0.02 | 0.63 | 1.38 | 0.76 | 2.54 | 0.32 | 0.91 | 1.39 | 0.72 | 0.81 | 0.33 | 0.53 | 1.5 | 0.41 | 0.47 | 3.02 | 0.07 | 1.31 | 0.07 | 0.17 | 0.38 | 0.32 | 0.12 | 4.76 | 0.02
| Analytical results for Pd in OREAS 45f (Indicative Value 59.3 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 39 |  | ppb
|  |  | 02: 5A*MS |  | 1
1
| 1 | 1 | 02: 73 |  | 1
| 1 | 2 | 02: 41 |  | 17
| 1 | 3 | 02: 69 |  | 16
| 1 | 4 | 02: 48 |  | 59.3333
| 1 | 5 | 02: 65 |  | 23
| 1 | 6 | 02: 60
| Mean |  | 02: 59.3333
| Median |  | 02: 62.5
| Std Dev. |  | 02: 12.4686
| Rel.Std.Dev. |  | 02: 21.01%
| PDM3 |  | 02: 6.661e-14%
Z-Score (Absolute) | bar: 02
| Analytical results for Pr in OREAS 45f (Certified Value 3.43 ppm) | [32 empty] | Certified
Pr | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab |  | 1
| No | No | 02: 02; 04: 12; 05: 13; 07: 15; 10: 18; 12: 23; 13: 24; 14: 27; 15: 33; 17: 38; 18: 39; 19: 40 |  | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*OES/MS; 15: 4A*MS; 17: 4A*MS; 18: 5A*OES; 19: 4A*MS |  | 2
3
| 1 | 1 | 02: 3.7235; 04: 3.42; 05: 3.37; 07: 3.33; 10: 4.91; 12: 3.3; 13: 3.4801; 14: 3.11; 15: 4.05; 17: 3.3267; 18: 1.522; 19: 3.1412 |  | 1
| 1 | 2 | 02: 3.7432; 04: 3.44; 05: 3.45; 07: 3.23; 10: 4.92; 12: 3.35; 13: 3.4824; 14: 3.04; 15: 4.07; 17: 3.53; 18: 1.444; 19: 3.1792 |  | #N/A
| 1 | 3 | 02: 3.6933; 04: 3.34; 05: 3.61; 07: 3.59; 10: 4.94; 12: 3.39; 13: 3.4633; 14: 3.13; 15: 3.93; 17: 3.3067; 18: 1.465; 19: 3.0766 |  | 16
| 1 | 4 | 02: 3.7981; 04: 3.48; 05: 3.19; 07: 3.39; 10: 4.46; 12: 3.41; 13: 3.512; 14: 3.02; 15: 3.03; 17: 3.43; 18: 1.597; 19: 2.8722 |  | 3.4311
| 1 | 5 | 02: 3.7324; 04: 3.36; 05: 3.49; 07: 3.19; 10: 4.16; 12: 3.28; 13: 3.5769; 14: 2.93; 15: 3.98; 17: 3.3133; 18: 1.316; 19: 3.2325 |  | 116
| 1 | 6 | 02: 3.6072; 04: 3.38; 05: 3.71; 07: 3.39; 10: 4.02; 12: 3.41; 13: 3.5844; 14: 3.02; 15: 4.03; 17: 3.3333; 18: 1.371; 19: 3.0912
| Mean |  | 02: 3.7163; 04: 3.4033; 05: 3.47; 07: 3.3533; 10: 4.5683; 12: 3.3567; 13: 3.5165; 14: 3.0417; 15: 3.8483; 17: 3.3733; 18: 1.4525; 19: 3.0988
| Median |  | 02: 3.7279; 04: 3.4; 05: 3.47; 07: 3.36; 10: 4.685; 12: 3.37; 13: 3.4972; 14: 3.03; 15: 4.005; 17: 3.33; 18: 1.4545; 19: 3.1162
| Std Dev. |  | 02: 0.063494; 04: 0.052789; 05: 0.18243; 07: 0.14222; 10: 0.41417; 12: 0.056451; 13: 0.052159; 14: 0.071949; 15: 0.4041; 17: 0.088969; 18: 0.10121; 19: 0.12493
| Rel.Std.Dev. |  | 02: 1.709%; 04: 1.551%; 05: 5.257%; 07: 4.241%; 10: 9.066%; 12: 1.682%; 13: 1.483%; 14: 2.365%; 15: 10.5%; 17: 2.637%; 18: 6.968%; 19: 4.031%
| PDM3 |  | 02: 8.313%; 04: -0.8082%; 05: 1.135%; 07: -2.265%; 10: 33.15%; 12: -2.168%; 13: 2.491%; 14: -11.35%; 15: 12.16%; 17: -1.683%; 18: -57.67%; 19: -9.683%
| Z-Score (Absolute) |  | 02: 1.06; 04: 0.05; 05: 0.26; 07: 0.11; 10: 3.81; 12: 0.1; 13: 0.41; 14: 1.12; 15: 1.49; 17: 0.05; 18: 6.25; 19: 0.93
| Analytical results for Pt in OREAS 45f (Indicative Value 41.5 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | 02: Lab; 04: Lab |  | 1
| No | No | 02: 10; 04: 39 |  | ppb
|  |  | 02: 4A*OES/MS; 04: 5A*MS |  | 1
1
| 1 | 1 | 02: 240.3956; 04: 47 |  | 1
| 1 | 2 | 02: 266.9316; 04: 39 |  | 10
| 1 | 3 | 02: 235.7272; 04: 44 |  | 16
| 1 | 4 | 02: 259.0778; 04: 42 |  | 41.5
| 1 | 5 | 02: 229.0949; 04: 39 |  | 24
| 1 | 6 | 02: 241.6839; 04: 38
| Mean |  | 02: 245.4852; 04: 41.5
| Median |  | 02: 241.0398; 04: 40.5
| Std Dev. |  | 02: 14.4839; 04: 3.5071
| Rel.Std.Dev. |  | 02: 5.9%; 04: 8.451%
| PDM3 |  | 02: 491.5%; 04: 0%
Z-Score (Absolute) | bar: 02 to 04
Results from laboratory 10 were removed due to their lack of accuracy.
| Analytical results for Rb in OREAS 45f (Certified Value 31.2 ppm) | [32 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 33.326 | 35.3 | 33.2 | 31.3 | 87.6108 | 31 | 32.23 | 27.9 | 22 | 37.6 | 31.4 | 32.7 | 45.8 | 30.64 | 31.6951 | 31.9 | 34.8 | 31 | 29.6 | 32.7 | 30.5 | 32.4 | 29.3 | 31.1 | 35 | 30.6033 | 28.915 | 30.749 |  |  |  | 1
| 1 | 2 | 33.52 | 35.7 | 32.2 | 32.3 | 83.4212 | 31.4 | 32.78 | 28.6 | 22 | 34.9 | 30 | 29.9 | 49.4 | 31.53 | 31.8545 | 32.1 | 33.6 | 29.6 | 30.9 | 30.8 | 30.7 | 31 | 29.6 | 31.1 | 35.4 | 32.3367 | 26.841 | 30.6087 |  |  |  | 2
| 1 | 3 | 33.168 | 33.5 | 33.6 | 31.6 | 84.3758 | 29.4 | 32.26 | 27.2 | 26 | 38.1 | 30.4 | 31.3 | 40.1 | 31.11 | 31.5429 | 32.3 | 33.4 | 29.6 | 29.4 | 30.7 | 30.1 | 30.2 | 25.1 | 31.5 | 33.2 | 30.17 | 28.149 | 31.4064 |  |  |  | 16
| 1 | 4 | 33.59 | 35 | 33.8 | 31.5 | 80.1066 | 31 | 31.62 | 28.9 | 24 | 36.9 | 31.8 | 31.7 | 46.5 | 31.46 | 31.8644 | 33.2 | 33.7 | 27.4 | 29 | 31.3 | 30.3 | 34.1 | 26.4 | 31.7 | 32.8 | 30.96 | 25.879 | 30.9755 |  |  |  | 31.2439
| 1 | 5 | 33.158 | 33.5 | 34.2 | 31.6 | 89.2839 | 29.6 | 32.66 | 26.9 | 21 | 37.1 | 31.2 | 32.9 | 41.6 | 30.94 | 31.6179 | 32.2 | 33.1 | 28.7 | 30.3 | 30.5 | 27.8 | 34 | 29.8 | 31.5 | 33.7 | 30.9433 | 26.768 | 31.3349 |  |  |  | 117
| 1 | 6 | 33.649 | 33 | 32.2 | 32 | 82.4017 | 30.4 | 33.66 | 24.3 | 23 | 35 | 29.6 | 31.8 | 42.3 | 30.68 | 31.6324 | 31.7 | 34.1 | 29.8 | 29 | 30.6 | 28.9 | 32.6 | 29.6 | 30.6 | 32.7 | 29.65 | 27.361 | 30.1197
| Mean |  | 33.4018 | 34.3333 | 33.2 | 31.7167 | 84.5333 | 30.4667 | 32.535 | 27.3 | 23 | 36.6 | 30.7333 | 31.7167 | 44.2833 | 31.06 | 31.7012 | 32.2333 | 33.7833 | 29.35 | 29.7 | 31.1 | 29.7167 | 32.3833 | 28.3 | 31.25 | 33.8 | 30.7772 | 27.3188 | 30.8657
| Median |  | 33.423 | 34.25 | 33.4 | 31.6 | 83.8985 | 30.7 | 32.46 | 27.55 | 22.5 | 37 | 30.8 | 31.75 | 44.05 | 31.025 | 31.6638 | 32.15 | 33.65 | 29.6 | 29.5 | 30.75 | 30.2 | 32.5 | 29.45 | 31.3 | 33.45 | 30.7733 | 27.101 | 30.8622
| Std Dev. |  | 0.21464 | 1.1325 | 0.83905 | 0.3656 | 3.3894 | 0.8165 | 0.68515 | 1.6601 | 1.7889 | 1.3446 | 0.8641 | 1.0815 | 3.5222 | 0.3791 | 0.13182 | 0.52026 | 0.59805 | 1.2062 | 0.75895 | 0.83187 | 1.1321 | 1.5677 | 2.0239 | 0.39875 | 1.1472 | 0.91241 | 1.0811 | 0.48182
| Rel.Std.Dev. |  | 0.6426% | 3.299% | 2.527% | 1.153% | 4.01% | 2.68% | 2.106% | 6.081% | 7.778% | 3.674% | 2.812% | 3.41% | 7.954% | 1.221% | 0.4158% | 1.614% | 1.77% | 4.11% | 2.555% | 2.675% | 3.81% | 4.841% | 7.151% | 1.276% | 3.394% | 2.965% | 3.957% | 1.561%
| PDM3 |  | 6.907% | 9.888% | 6.261% | 1.513% | 170.6% | -2.488% | 4.132% | -12.62% | -26.39% | 17.14% | -1.634% | 1.513% | 41.73% | -0.5885% | 1.464% | 3.167% | 8.128% | -6.062% | -4.941% | -0.4605% | -4.888% | 3.647% | -9.422% | 0.01959% | 8.181% | -1.494% | -12.56% | -1.21%
| Z-Score (Absolute) |  | 0.75 | 1.11 | 0.67 | 0.09 | 20.54 | 0.39 | 0.41 | 1.62 | 3.28 | 1.98 | 0.29 | 0.09 | 4.96 | 0.16 | 0.09 | 0.29 | 0.89 | 0.82 | 0.69 | 0.15 | 0.68 | 0.35 | 1.23 | 0.09 | 0.9 | 0.27 | 1.61 | 0.24
| Analytical results for Re in OREAS 45f (Indicative Value < 0.002 ppm) | [32 empty] | Indicative
Re | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab |  | 1
| No | No | 02: 07; 04: 12; 05: 13; 07: 14; 10: 19; 12: 20; 13: 21; 14: 24; 15: 28; 17: 29; 18: 30; 19: 31; 20: 35; 21: 36; 23: 38; 24: 39 |  | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*MS; 17: 4A*MS; 18: 4A*OES/MS; 19: 4A*MS; 20: 4A*MS; 21: 4A*MS; 23: 4A*MS; 24: 5A*MS |  | 3
3
| 1 | 1 | 02: < 0.05; 04: < 0.1; 05: < 0.001; 07: < 0.002; 10: < 0.1; 12: < 0.005; 13: < 0.005; 14: < 0.002; 15: < 0.002; 17: < 0.002; 18: < 0.002; 19: < 0.002; 20: < 0.002; 21: 0.003; 23: < 0.002; 24: < 0.002 |  | 1
| 1 | 2 | 02: < 0.05; 04: < 0.1; 05: 0.003; 07: < 0.002; 10: < 0.1; 12: 0.008; 13: < 0.005; 14: < 0.002; 15: < 0.002; 17: < 0.002; 18: 0.003; 19: < 0.002; 20: < 0.002; 21: 0.003; 23: < 0.002; 24: < 0.002 |  | 11
| 1 | 3 | 02: < 0.05; 04: < 0.1; 05: < 0.001; 07: < 0.002; 10: < 0.1; 12: < 0.005; 13: < 0.005; 14: < 0.002; 15: < 0.002; 17: < 0.002; 18: 0.002; 19: < 0.002; 20: < 0.002; 21: 0.002; 23: < 0.002; 24: < 0.002 |  | 16
| 1 | 4 | 02: < 0.05; 04: < 0.1; 05: 0.001; 07: < 0.002; 10: < 0.1; 12: 0.008; 13: < 0.005; 14: < 0.002; 15: < 0.002; 17: < 0.002; 18: < 0.002; 19: < 0.002; 20: < 0.002; 21: 0.002; 23: < 0.002; 24: < 0.002 |  | < 0.002
| 1 | 5 | 02: < 0.05; 04: < 0.1; 05: < 0.001; 07: < 0.002; 10: < 0.1; 12: < 0.005; 13: < 0.005; 14: < 0.002; 15: < 0.002; 17: < 0.002; 18: < 0.002; 19: < 0.002; 20: < 0.002; 21: 0.003; 23: < 0.002; 24: < 0.002 |  | 25
| 1 | 6 | 02: < 0.05; 04: < 0.1; 05: < 0.001; 07: < 0.002; 10: < 0.1; 12: < 0.005; 13: < 0.005; 14: < 0.002; 15: < 0.002; 17: < 0.002; 18: < 0.002; 19: < 0.002; 20: < 0.002; 21: 0.003; 23: < 0.002; 24: < 0.002
| Mean |  | 05: 0.002; 12: 0.008; 18: 0.0025; 21: 0.0026667
| Median |  | 05: 0.002; 12: 0.008; 18: 0.0025; 21: 0.003
| Std Dev. |  | 05: 0.0014142; 12: 0; 18: 0.00070711; 21: 0.0005164
| Rel.Std.Dev. |  | 05: 70.71%; 12: 0%; 18: 28.28%; 21: 19.36%
PDM3
Z-Score (Absolute) | bar: 02 to 24
| Analytical results for Rh in OREAS 45f (Indicative Value < 5 ppb) | [32 empty] | Indicative
Rh | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 39 |  | ppb
|  |  | 02: 5A*MS |  | 2
2
| 1 | 1 | 02: < 5 |  | 1
| 1 | 2 | 02: < 5 |  | 12
| 1 | 3 | 02: < 5 |  | 16
| 1 | 4 | 02: < 5 |  | < 5
| 1 | 5 | 02: < 5 |  | 21
| 1 | 6 | 02: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for Ru in OREAS 45f (Indicative Value < 10 ppb) | [32 empty] | Indicative
Ru | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 39 |  | ppb
|  |  | 02: 5A*MS |  | 1
1
| 1 | 1 | 02: < 5 |  | 1
| 1 | 2 | 02: < 5 |  | 13
| 1 | 3 | 02: 7 |  | 16
| 1 | 4 | 02: < 5 |  | < 10
| 1 | 5 | 02: < 5 |  | 22
| 1 | 6 | 02: 5
| Mean |  | 02: 6
| Median |  | 02: 6
| Std Dev. |  | 02: 1.4142
| Rel.Std.Dev. |  | 02: 23.57%
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for S in OREAS 45f (Certified Value 0.029 wt.%) | [32 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 04 | 05 | 07 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.03 | 0.03 | 0.028 | 0.035 | 0.03 | 0.04 | 0.028 | 0.03 | 0.0235 | 0.03 | < 0.1 | 0.03 | 0.0278 | 0.026776 | 0.0312 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0285 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0285 | 0.0249 | 0.024194 |  |  |  | 1
| 1 | 2 | 0.03 | 0.03 | 0.027 | 0.03 | 0.03 | 0.04 | 0.0287 | 0.02 | 0.0265 | 0.025 | < 0.1 | 0.03 | 0.0277 | 0.027855 | 0.0313 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.028 | 0.03 | 0.03 | 0.03 | 0.03 | 0.028267 | 0.0246 | 0.025041 |  |  |  | 28
| 1 | 3 | 0.03 | 0.03 | 0.028 | 0.035 | 0.03 | 0.04 | 0.0281 | 0.02 | 0.0245 | 0.03 | < 0.1 | 0.03 | 0.0275 | 0.027235 | 0.0311 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.028 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0276 | 0.0271 | 0.038006 |  |  |  | 16
| 1 | 4 | 0.04 | 0.03 | 0.027 | 0.03 | 0.03 | 0.04 | 0.0278 | 0.03 | 0.018 | 0.03 | < 0.1 | 0.03 | 0.0274 | 0.027585 | 0.0304 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0275 | 0.03 | 0.03 | 0.03 | 0.03 | 0.028467 | 0.0247 | 0.024641 |  |  |  | 0.029177
| 1 | 5 | 0.03 | 0.03 | 0.027 | 0.03 | 0.03 | 0.04 | 0.0263 | 0.03 | 0.023 | 0.03 | < 0.1 | 0.03 | 0.0274 | 0.027366 | 0.0316 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.029 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027533 | 0.0254 | 0.029368 |  |  |  | 118
| 1 | 6 | 0.04 | 0.03 | 0.027 | 0.03 | 0.03 | 0.04 | 0.0285 | 0.02 | 0.0225 | 0.025 | < 0.1 | 0.03 | 0.0276 | 0.026874 | 0.031 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.0285 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027633 | 0.0262 | 0.03185
| Mean |  | 0.033333 | 0.03 | 0.027333 | 0.031667 | 0.03 | 0.04 | 0.0279 | 0.025 | 0.023 | 0.028333 |  | 0.03 | 0.027567 | 0.027282 | 0.0311 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02825 | 0.03 | 0.03 | 0.03 | 0.03 | 0.028 | 0.025483 | 0.02885
| Median |  | 0.03 | 0.03 | 0.027 | 0.03 | 0.03 | 0.04 | 0.02805 | 0.025 | 0.02325 | 0.03 |  | 0.03 | 0.02755 | 0.0273 | 0.03115 | 0.04 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02825 | 0.03 | 0.03 | 0.03 | 0.03 | 0.02795 | 0.02515 | 0.027204
| Std Dev. |  | 0.005164 | 0 | 0.0005164 | 0.002582 | 0 | 0 | 0.00085088 | 0.0054772 | 0.0028284 | 0.002582 |  | 0 | 0.0001633 | 0.00041269 | 0.0004 | 0 | 0 | 0 | 0 | 0 | 0.0005244 | 0 | 0 | 0 | 0 | 0.0004585 | 0.00098675 | 0.0054224
| Rel.Std.Dev. |  | 15.49% | 0% | 1.889% | 8.154% | 0% | 0% | 3.05% | 21.91% | 12.3% | 9.113% |  | 0% | 0.5924% | 1.513% | 1.286% | 0% | 0% | 0% | 0% | 0% | 1.856% | 0% | 0% | 0% | 0% | 1.637% | 3.872% | 18.8%
| PDM3 |  | 14.24% | 2.82% | -6.32% | 8.532% | 2.82% | 37.09% | -4.378% | -14.32% | -21.17% | -2.892% |  | 2.82% | -5.52% | -6.496% | 6.59% | 37.09% | 2.82% | 2.82% | 2.82% | 2.82% | -3.178% | 2.82% | 2.82% | 2.82% | 2.82% | -4.035% | -12.66% | -1.121%
| Z-Score (Absolute) |  | 1.35 | 0 | 1.08 | 0.67 | 0 | 4.05 | 0.85 | 2.02 | 2.83 | 0.67 | 8.09 | 0 | 0.98 | 1.1 | 0.45 | 4.05 | 0 | 0 | 0 | 0 | 0.71 | 0 | 0 | 0 | 0 | 0.81 | 1.83 | 0.47
| Analytical results for Sb in OREAS 45f (Certified Value 0.64 ppm) | [32 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 0.67 | 0.72 | 0.7 | 0.82169 | 0.4 | 0.68 | 0.6 | 0.9 | < 5 | 0.6 | 0.4 | 0.6 | 0.7 | 0.53 | 0.75465 | 0.641 | 0.66 | 0.6 | 0.6 | 0.68 | 0.6 | 0.7 | 0.68 | 0.59 | 0.89 | 6 | 0.56 | 0.462 | 0.5966 |  |  | 1
| 1 | 2 | 0.68 | 0.69 | 0.7 | 0.79089 | 0.4 | 0.6 | 0.65 | 0.9 | < 5 | 0.6 | 0.5 | 0.5 | 0.7 | 0.62 | 0.7113 | 0.629 | 0.66 | 0.62 | 0.56 | 0.71 | 0.55 | 0.5 | 0.68 | 0.58 | 0.77 | 5 | 0.60333 | 0.293 | 0.6306 |  |  | 26
| 1 | 3 | 0.6 | 0.84 | 0.6 | 0.83904 | 0.6 | 0.65 | 0.61 | 0.8 | < 5 | 0.6 | 0.5 | 0.6 | 0.7 | 0.5 | 0.72543 | 0.653 | 0.67 | 0.56 | 0.59 | 0.65 | 0.57 | 0.6 | 0.69 | 0.64 | 0.76 | 5 | 0.57333 | 0.431 | 0.6122 |  |  | 16
| 1 | 4 | 0.65 | 0.66 | 0.6 | 0.87582 | 0.6 | 0.69 | 0.61 | 0.9 | < 5 | 0.6 | 0.6 | 0.7 | 0.8 | 0.6 | 0.77372 | 0.581 | 0.67 | 0.54 | 0.63 | 0.72 | 0.55 | 0.5 | 0.7 | 0.66 | 0.85 | < 5 | 0.59 | 0.304 | 0.6267 |  |  | 0.63635
| 1 | 5 | 0.63 | 0.69 | 0.6 | 0.86783 | 0.4 | 0.69 | 0.61 | 0.6 | < 5 | 0.6 | 0.6 | 0.6 | 0.7 | 0.64 | 0.80273 | 0.634 | 0.68 | 0.62 | 0.53 | 0.65 | 0.55 | 0.6 | 0.68 | 0.62 | 0.85 | 5 | 0.56333 | 0.349 | 0.6132 |  |  | 119
| 1 | 6 | 0.62 | 0.67 | 0.7 | 0.95644 | 0.4 | 0.53 | 0.58 | 0.9 | < 5 | 0.6 | 0.5 | 0.5 | 0.6 | 0.53 | 0.79402 | 0.6 | 0.64 | 0.61 | 0.64 | 0.65 | 0.59 | 0.7 | 0.71 | 0.6 | 0.8 | 5 | 0.55667 | 0.386 | 0.6163
| Mean |  | 0.64167 | 0.71167 | 0.65 | 0.85862 | 0.46667 | 0.64 | 0.61 | 0.83333 |  | 0.6 | 0.51667 | 0.58333 | 0.7 | 0.57 | 0.76031 | 0.623 | 0.66333 | 0.59167 | 0.59167 | 0.67667 | 0.56833 | 0.6 | 0.69 | 0.615 | 0.82 | 5.2 | 0.57444 | 0.37083 | 0.61593
| Median |  | 0.64 | 0.69 | 0.65 | 0.85344 | 0.4 | 0.665 | 0.61 | 0.9 |  | 0.6 | 0.5 | 0.6 | 0.7 | 0.565 | 0.76419 | 0.6315 | 0.665 | 0.605 | 0.595 | 0.665 | 0.56 | 0.6 | 0.685 | 0.61 | 0.825 | 5 | 0.56833 | 0.3675 | 0.61475
| Std Dev. |  | 0.030605 | 0.066156 | 0.054772 | 0.057069 | 0.10328 | 0.063875 | 0.022804 | 0.12111 |  | 0 | 0.075277 | 0.075277 | 0.063246 | 0.057271 | 0.036787 | 0.027107 | 0.013663 | 0.033714 | 0.041673 | 0.032042 | 0.022286 | 0.089443 | 0.012649 | 0.030822 | 0.051381 | 0.44721 | 0.018579 | 0.068063 | 0.012053
| Rel.Std.Dev. |  | 4.77% | 9.296% | 8.427% | 6.647% | 22.13% | 9.98% | 3.738% | 14.53% |  | 0% | 14.57% | 12.9% | 9.035% | 10.05% | 4.838% | 4.351% | 2.06% | 5.698% | 7.043% | 4.735% | 3.921% | 14.91% | 1.833% | 5.012% | 6.266% | 8.6% | 3.234% | 18.35% | 1.957%
| PDM3 |  | 0.8348% | 11.83% | 2.144% | 34.93% | -26.67% | 0.5729% | -4.141% | 30.95% |  | -5.713% | -18.81% | -8.332% | 10% | -10.43% | 19.48% | -2.099% | 4.24% | -7.022% | -7.022% | 6.335% | -10.69% | -5.713% | 8.43% | -3.356% | 28.86% | 717.2% | -9.729% | -41.73% | -3.209%
| Z-Score (Absolute) |  | 0.02 | 0.97 | NA | 2.95 | NA | 0 | 0.4 | NA | 25.08 | NA | NA | NA | 0.81 | 0.94 | 1.62 | 0.23 | 0.31 | 0.65 | 0.65 | 0.49 | 0.97 | NA | 0.67 | 0.34 | 2.43 | NA | 0.88 | 3.63 | 0.32
Results from laboratories 7, 12, 15, 18, 19, 20, 33 & 36 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Sc in OREAS 45f (Certified Value 36.3 ppm) | [32 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 35.826 | 36.6 | 38.6 | 35 | 51.7979 | 37 | 37.1 | 35.7 | 37 | 33 | 45.6 | 36 | 36 | 37.2 | 31.7929 | 35.6 | 48 | 40.1 | 35.6 | 33 | 40.8 | 40.2 | 34.1 | 34.6 | 39.8 | 36.7 | 36.1 | 31.408 | 35.7748 |  |  | 1
| 1 | 2 | 36.162 | 35.5 | 38.5 | 35 | 50.6936 | 38 | 40.7 | 35.7 | 37.9 | 19 | 40.1 | 34 | 39 | 36.8 | 32.16 | 34.6 | 45.7 | 37 | 37.4 | 37.9 | 39.6 | 32.7 | 36.8 | 34.9 | 39.3 | 36.2 | 38.2667 | 31.941 | 35.7524 |  |  | #N/A
| 1 | 3 | 35.541 | 33.6 | 39.2 | 35 | 48.0506 | 38 | 37.1 | 33.6 | 38.4 | 30 | 44.3 | 35 | 33 | 37.1 | 31.8207 | 36.4 | 46.3 | 39.5 | 36.4 | 38 | 39.7 | 38.3 | 33.6 | 36.4 | 40.1 | 34.9 | 35.6333 | 32.277 | 36.7372 |  |  | 16
| 1 | 4 | 35.987 | 34.4 | 40.3 | 34 | 49.8937 | 38 | 37.6 | 35.6 | 37.7 | 24 | 52.2 | 35 | 37 | 37.1 | 31.6535 | 35.9 | 47.1 | 36.1 | 37.8 | 38.8 | 39.5 | 36.6 | 31.4 | 36.3 | 39.4 | 36.5 | 36.6333 | 31.837 | 35.3603 |  |  | 36.2753
| 1 | 5 | 36.366 | 35.8 | 41 | 35 | 47.4808 | 37 | 36.5 | 35.4 | 37.4 | 29 | 48.2 | 36 | 34 | 37.3 | 32.1933 | 35.8 | 47 | 38.4 | 35.9 | 37.7 | 34.4 | 37.1 | 36.2 | 35.6 | 40.5 | 34.9 | 35.9667 | 31.964 | 36.8282 |  |  | 120
| 1 | 6 | 35.775 | 33.9 | 39.5 | 35 | 49.8122 | 37 | 37.4 | 33.4 | 38 | 24 | 45.8 | 36 | 35 | 37.2 | 31.8146 | 35.5 | 41.3 | 38.3 | 34.6 | 38.1 | 37.5 | 42.5 | 34.4 | 34.1 | 39.5 | 35.9 | 35.2 | 31.351 | 33.7776
| Mean |  | 35.9428 | 34.9667 | 39.5167 | 34.8333 | 49.6215 | 37.5 | 37.7333 | 34.9 | 37.7333 | 26.5 | 46.0333 | 35.3333 | 35.6667 | 37.1167 | 31.9058 | 35.6333 | 45.9 | 38.2333 | 36.2833 | 37.25 | 38.5833 | 37.9 | 34.4167 | 35.3167 | 39.7667 | 35.85 | 36.3 | 31.7963 | 35.7051
| Median |  | 35.9065 | 34.95 | 39.35 | 35 | 49.8529 | 37.5 | 37.25 | 35.5 | 37.8 | 26.5 | 45.7 | 35.5 | 35.5 | 37.15 | 31.8177 | 35.7 | 46.65 | 38.35 | 36.15 | 37.95 | 39.55 | 37.7 | 34.25 | 35.25 | 39.65 | 36.05 | 36.0333 | 31.889 | 35.7636
| Std Dev. |  | 0.29406 | 1.181 | 0.9786 | 0.40825 | 1.6154 | 0.54772 | 1.5002 | 1.0918 | 0.48854 | 5.0892 | 4.0292 | 0.8165 | 2.1602 | 0.17224 | 0.21873 | 0.59554 | 2.3841 | 1.4962 | 1.1839 | 2.1154 | 2.3112 | 3.346 | 1.9354 | 0.93684 | 0.46332 | 0.78422 | 1.0756 | 0.35521 | 1.1107
| Rel.Std.Dev. |  | 0.8181% | 3.377% | 2.476% | 1.172% | 3.255% | 1.461% | 3.976% | 3.128% | 1.295% | 19.2% | 8.753% | 2.311% | 6.057% | 0.4641% | 0.6855% | 1.671% | 5.194% | 3.913% | 3.263% | 5.679% | 5.99% | 8.829% | 5.623% | 2.653% | 1.165% | 2.188% | 2.963% | 1.117% | 3.111%
| PDM3 |  | -0.9166% | -3.608% | 8.935% | -3.975% | 36.79% | 3.376% | 4.019% | -3.791% | 4.019% | -26.95% | 26.9% | -2.597% | -1.678% | 2.319% | -12.05% | -1.77% | 26.53% | 5.398% | 0.02204% | 2.687% | 6.362% | 4.479% | -5.124% | -2.643% | 9.625% | -1.173% | 0.06799% | -12.35% | -1.572%
| Z-Score (Absolute) |  | 0.16 | 0.61 | 1.5 | 0.67 | 6.2 | 0.57 | 0.67 | 0.64 | 0.67 | 4.55 | 4.53 | 0.44 | 0.29 | 0.39 | 2.04 | 0.3 | 4.47 | 0.91 | 0 | 0.45 | 1.07 | 0.75 | 0.87 | 0.45 | 1.62 | 0.2 | 0.01 | 2.09 | 0.27
| Analytical results for Se in OREAS 45f (Certified Value 2.26 ppm) | [32 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | [5 empty] | 2
2
| 1 | 1 | < 2 | 2 | 3 | 1.8525 | < 5 | 2.3 | 2.4 | 2.03 | 3 | < 5 | 3 | < 1 | 2.4 | 2.2304 | < 2 | 2 | 2.3 | 1.8 | 4 | 2.6 | 2 | < 2 | 2.5 | 3.1 | 2.4333 | 1.894 | [5 empty] | 1
| 1 | 2 | < 2 | 2 | 3 | 1.8164 | < 5 | 1.3 | 2.5 | 2.94 | 2 | < 5 | 2 | < 1 | 2.1 | 2.1279 | 2.2 | 2 | 2.2 | 2.2 | 4 | 2.6 | 2 | < 2 | 2.8 | 2.7 | 2.5333 | 1.705 | [5 empty] | 8
| 1 | 3 | < 2 | 2 | 2 | 1.7779 | < 5 | 2.9 | 2.1 | 2.95 | 3 | < 5 | 3 | < 1 | 2 | 2.1749 | 2.1 | 2 | 2.4 | 1.7 | 3 | 2.5 | 2 | < 2 | 2.7 | 2.7 | 2.5 | 1.802 | [5 empty] | 16
| 1 | 4 | < 2 | 3 | 3 | 1.8239 | < 5 | 1.4 | 2.1 | 2.93 | 3 | < 5 | 2 | < 1 | 1.9 | 1.9313 | 2.4 | 2 | 2.3 | 2.5 | 5 | 2.6 | 2 | < 2 | 2.7 | 2.7 | 2.4 | 1.646 | [5 empty] | 2.2628
| 1 | 5 | < 2 | 2 | 2 | 1.847 | < 5 | 1.8 | 2.2 | 2.67 | 2 | < 5 | 2 | < 1 | 2.1 | 1.8179 | 2.3 | 2 | 2.2 | 1.9 | 3 | 2.5 | 2 | < 2 | 2.7 | 2.8 | 2.4 | 1.722 | [5 empty] | 121
| 1 | 6 | < 2 | 3 | 3 | 1.8205 | < 5 | 4.9 | 2.2 | 2.63 | 2 | < 5 | 3 | < 1 | 2.5 | 1.8138 | 2.3 | 2 | 2.5 | 2.1 | 4 | 2.3 | 2 | < 2 | 2.6 | 2.5 | 2.3333 | 1.746
| Mean |  |  | 2.3333 | 2.6667 | 1.823 |  | 2.4333 | 2.25 | 2.6917 | 2.5 |  | 2.5 |  | 2.1667 | 2.016 | 2.26 | 2 | 2.3167 | 2.0333 | 3.8333 | 2.5167 | 2 |  | 2.6667 | 2.75 | 2.4333 | 1.7525
| Median |  |  | 2 | 3 | 1.8222 |  | 2.05 | 2.2 | 2.8 | 2.5 |  | 2.5 |  | 2.1 | 2.0296 | 2.3 | 2 | 2.3 | 2 | 4 | 2.55 | 2 |  | 2.7 | 2.7 | 2.4167 | 1.734
| Std Dev. |  |  | 0.5164 | 0.5164 | 0.026581 |  | 1.3471 | 0.16432 | 0.3542 | 0.54772 |  | 0.54772 |  | 0.23381 | 0.18493 | 0.11402 | 0 | 0.1169 | 0.29439 | 0.75277 | 0.1169 | 0 |  | 0.10328 | 0.19748 | 0.07303 | 0.086074
| Rel.Std.Dev. |  |  | 22.13% | 19.36% | 1.458% |  | 55.36% | 7.303% | 13.16% | 21.91% |  | 21.91% |  | 10.79% | 9.173% | 5.045% | 0% | 5.046% | 14.48% | 19.64% | 4.645% | 0% |  | 3.873% | 7.181% | 3.001% | 4.911%
| PDM3 |  |  | 3.118% | 17.85% | -19.43% |  | 7.537% | -0.5647% | 18.95% | 10.48% |  | 10.48% |  | -4.247% | -10.9% | -0.1228% | -11.61% | 2.382% | -10.14% | 69.41% | 11.22% | -11.61% |  | 17.85% | 21.53% | 7.537% | -22.55%
| Z-Score (Absolute) |  | 3.37 | NA | NA | 1.15 | 0.67 | 0.49 | 0 | 1.19 | NA | 0.67 | NA | 4.72 | 0.22 | 0.63 | 0.54 | NA | 0.18 | 0.58 | NA | 0.72 | NA | 3.37 | 1.12 | 1.35 | 0.49 | 1.34
Results from laboratories 5, 7, 18, 20, 27, 30, & 33 were removed due to their 1 ppm reading resolution.
| Analytical results for Sm in OREAS 45f (Certified Value 2.49 ppm) | [32 empty] | Certified
Sm | Round | Replicate | 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab |  |  | 1
| No | No | 04: 02; 05: 12; 07: 13; 10: 15; 12: 18; 13: 23; 14: 24; 15: 27; 17: 33; 18: 38; 19: 39; 20: 40 |  |  | ppm
|  |  | 04: 4A*MS; 05: 4A*MS; 07: 4A*OES/MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*OES/MS; 17: 4A*MS; 18: 4A*MS; 19: 5A*OES; 20: 4A*MS |  |  | 2
3
| 1 | 1 | 04: 2.7296; 05: 2.5; 07: 2.48; 10: 2.24; 12: 3.6; 13: 2.52; 14: 2.4925; 15: 2.5; 17: 2.5; 18: 2.4033; 19: 0.657; 20: 2.3316 |  |  | 1
| 1 | 2 | 04: 2.6652; 05: 2.7; 07: 2.53; 10: 2.18; 12: 3.4; 13: 2.4; 14: 2.6208; 15: 2.6; 17: 2.6; 18: 2.5367; 19: 0.869; 20: 2.4374 |  |  | #N/A
| 1 | 3 | 04: 2.8364; 05: 2.4; 07: 2.68; 10: 2.56; 12: 3.7; 13: 2.41; 14: 2.5059; 15: 2.4; 17: 2.5; 18: 2.38; 19: 0.54; 20: 2.2931 |  |  | 16
| 1 | 4 | 04: 2.7534; 05: 2.6; 07: 2.29; 10: 2.44; 12: 3.6; 13: 2.45; 14: 2.4892; 15: 2.5; 17: 2.5; 18: 2.42; 19: 0.786; 20: 2.1821 |  |  | 2.4874
| 1 | 5 | 04: 2.7389; 05: 2.45; 07: 2.61; 10: 2.15; 12: 3.2; 13: 2.58; 14: 2.6132; 15: 2.4; 17: 2.6; 18: 2.38; 19: 0.732; 20: 2.4053 |  |  | 122
| 1 | 6 | 04: 2.6155; 05: 2.65; 07: 2.56; 10: 2.42; 12: 3.1; 13: 2.48; 14: 2.547; 15: 2.3; 17: 2.6; 18: 2.4; 19: 0.61; 20: 2.3282
| Mean |  | 04: 2.7232; 05: 2.55; 07: 2.525; 10: 2.3317; 12: 3.4333; 13: 2.4733; 14: 2.5448; 15: 2.45; 17: 2.55; 18: 2.42; 19: 0.699; 20: 2.3296
| Median |  | 04: 2.7342; 05: 2.55; 07: 2.545; 10: 2.33; 12: 3.5; 13: 2.465; 14: 2.5265; 15: 2.45; 17: 2.55; 18: 2.4017; 19: 0.6945; 20: 2.3299
| Std Dev. |  | 04: 0.076128; 05: 0.11832; 07: 0.13398; 10: 0.16497; 12: 0.24221; 13: 0.068605; 14: 0.05966; 15: 0.10488; 17: 0.054772; 18: 0.059142; 19: 0.12044; 20: 0.08994
| Rel.Std.Dev. |  | 04: 2.796%; 05: 4.64%; 07: 5.306%; 10: 7.075%; 12: 7.055%; 13: 2.774%; 14: 2.344%; 15: 4.281%; 17: 2.148%; 18: 2.444%; 19: 17.23%; 20: 3.861%
| PDM3 |  | 04: 9.478%; 05: 2.516%; 07: 1.511%; 10: -6.262%; 12: 38.03%; 13: -0.5664%; 14: 2.305%; 15: -1.504%; 17: 2.516%; 18: -2.71%; 19: -71.9%; 20: -6.344%
| Z-Score (Absolute) |  | 04: 2.32; 05: 0.53; 07: 0.27; 10: 1.74; 12: 9.69; 13: 0.27; 14: 0.47; 15: 0.51; 17: 0.53; 18: 0.82; 19: 18.67; 20: 1.76
| Analytical results for Sn in OREAS 45f (Certified Value 2.85 ppm) | [32 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.8 | 2.8 | 3 | 4.8067 | 3 | 2.8 | 2.7 | 2.7 | < 10 | 3.5 | 1.7 | 2.9 | 2.7 | 2.9 | 2.7206 | 2.82 | 4.1 | 3.16 | 2.67 | 3.2 | 2.92 | 2.9 | 4.6 | 2.75 | 3.1 | < 10 | 2.8333 | 2.755 | 2.9213 |  |  | 1
| 1 | 2 | 3.1 | 2.9 | 3.4 | 5.0605 | 3 | 3 | 2.61 | 2.9 | < 10 | 2.9 | 1.6 | 2.6 | 3.1 | 3 | 2.8098 | 2.79 | 3.8 | 2.57 | 2.78 | 2.9 | 2.88 | 3 | 3.7 | 2.76 | 3.1 | < 10 | 3 | 2.275 | 2.9649 |  |  | 3
| 1 | 3 | 2.5 | 3 | 3.2 | 4.747 | 3 | 2.8 | 2.6 | 2.6 | < 10 | 2.9 | 1.7 | 2.9 | 2.8 | 3 | 2.8389 | 2.71 | 3.7 | 2.54 | 2.56 | 3 | 2.73 | 2.7 | 3.7 | 2.82 | 3.2 | < 10 | 2.7667 | 2.488 | 3.0605 |  |  | 16
| 1 | 4 | 3 | 3 | 2.9 | 5.236 | 3 | 3 | 2.66 | 2.8 | < 10 | 3 | 1.6 | 2.9 | 2.8 | 2.9 | 2.7212 | 2.75 | 4 | 2.57 | 2.72 | 3.1 | 2.78 | 2.5 | 4.2 | 2.83 | 2.9 | < 10 | 2.8 | 2.223 | 2.9777 |  |  | 2.8506
| 1 | 5 | 2.9 | 3.1 | 2.9 | 4.8936 | 3 | 2.9 | 2.57 | 2.4 | < 10 | 2.8 | 1.8 | 3.1 | 2.8 | 2.9 | 2.6868 | 2.78 | 4.7 | 2.81 | 2.52 | 3 | 2.83 | 2.8 | 4.2 | 2.84 | 3.2 | < 10 | 2.8667 | 2.351 | 2.9892 |  |  | 123
| 1 | 6 | 2.9 | 3 | 3.4 | 5.0076 | 3 | 2.9 | 2.51 | 2.8 | < 10 | 2.9 | 1.5 | 2.9 | 2.8 | 2.7 | 2.7861 | 2.76 | 3.4 | 2.99 | 2.8 | 3.1 | 2.65 | 3 | 4.3 | 2.72 | 2.9 | < 10 | 2.7667 | 2.466 | 2.8848
| Mean |  | 2.8667 | 2.9667 | 3.1333 | 4.9586 | 3 | 2.9 | 2.6083 | 2.7 |  | 3 | 1.65 | 2.8833 | 2.8333 | 2.9 | 2.7606 | 2.7683 | 3.95 | 2.7733 | 2.675 | 3.05 | 2.7983 | 2.8167 | 4.1167 | 2.7867 | 3.0667 |  | 2.8389 | 2.4263 | 2.9664
| Median |  | 2.9 | 3 | 3.1 | 4.9506 | 3 | 2.9 | 2.605 | 2.75 |  | 2.9 | 1.65 | 2.9 | 2.8 | 2.9 | 2.7536 | 2.77 | 3.9 | 2.69 | 2.695 | 3.05 | 2.805 | 2.85 | 4.2 | 2.79 | 3.1 |  | 2.8167 | 2.4085 | 2.9713
| Std Dev. |  | 0.20656 | 0.10328 | 0.23381 | 0.17988 | 0 | 0.089443 | 0.066758 | 0.17889 |  | 0.25298 | 0.10488 | 0.16021 | 0.13663 | 0.10954 | 0.059642 | 0.037639 | 0.44159 | 0.25882 | 0.11485 | 0.10488 | 0.099482 | 0.19408 | 0.35449 | 0.049666 | 0.13663 |  | 0.087981 | 0.19148 | 0.060328
| Rel.Std.Dev. |  | 7.206% | 3.481% | 7.462% | 3.628% | 0% | 3.084% | 2.559% | 6.625% |  | 8.433% | 6.356% | 5.556% | 4.822% | 3.777% | 2.161% | 1.36% | 11.18% | 9.332% | 4.293% | 3.439% | 3.555% | 6.89% | 8.611% | 1.782% | 4.455% |  | 3.099% | 7.892% | 2.034%
| PDM3 |  | 0.5638% | 4.072% | 9.919% | 73.95% | 5.241% | 1.733% | -8.499% | -5.283% |  | 5.241% | -42.12% | 1.148% | -0.6056% | 1.733% | -3.159% | -2.886% | 38.57% | -2.71% | -6.16% | 6.995% | -1.833% | -1.19% | 44.41% | -2.243% | 7.58% |  | -0.4107% | -14.88% | 4.062%
| Z-Score (Absolute) |  | 0.05 | 0.52 | 1.46 | 11.74 | NA | 0.14 | 1.5 | 0.99 | 11.97 | 0.7 | 6.9 | 0.05 | 0.23 | 0.14 | 0.64 | 0.6 | 6.05 | 0.57 | 1.13 | 0.99 | 0.43 | 0.33 | 6.99 | 0.5 | 1.08 | 11.97 | 0.2 | 2.53 | 0.51
Results from laboratory 12 were removed due to their 1 ppm or greater reading resolution.
| Analytical results for Sr in OREAS 45f (Certified Value 25.1 ppm) | [32 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*MS |  | 1
2
| 1 | 1 | 24.593 | 26.6 | 25.8 | 23.8 | 37.7017 | 23.5 | 26 | 24 | 24.6 | 71 | 36 | 25 | 27 | 29 | 25.47 | 24.789 | 23.9 | 29.7 | 25.8 | 25.8 | 26.3 | 26.3 | 24.9 | 23.9 | 25.1 | 25.2 | 25 | 23.9867 | 26.322 | 24.1849 |  | 1
| 1 | 2 | 24.221 | 26.6 | 24.9 | 24.8 | 38.2393 | 24.5 | 28.6 | 23.9 | 23.7 | 60 | 32.5 | 24.5 | 26 | 26 | 25.29 | 25.2598 | 23.6 | 25.1 | 25.6 | 27.6 | 25.2 | 25.5 | 24.3 | 24.5 | 25.1 | 24.8 | 24.6 | 25.2667 | 26.827 | 23.9853 |  | #N/A
| 1 | 3 | 24.211 | 25.4 | 25.3 | 23.6 | 33.057 | 23 | 26.4 | 23 | 25.5 | 69 | 35.7 | 23.5 | 27 | 27 | 25.34 | 25.0491 | 23.8 | 26.8 | 25.8 | 26 | 25.2 | 25.8 | 23.7 | 23 | 25.9 | 24.9 | 24.3 | 23.94 | 26.598 | 24.8077 |  | 16
| 1 | 4 | 24.22 | 26 | 25.3 | 24 | 33.2333 | 24 | 27.1 | 24 | 24.9 | 70 | 38.3 | 24 | 26 | 24 | 25.49 | 25.2809 | 23.5 | 28.1 | 25 | 24.3 | 25.7 | 25.7 | 26.7 | 22.6 | 26.6 | 24.5 | 24.6 | 24.32 | 26.208 | 24.2805 |  | 25.1016
| 1 | 5 | 24.75 | 26.2 | 26.1 | 23.6 | 34.9285 | 23.5 | 25.9 | 23.9 | 23.3 | 68 | 35.8 | 24 | 27 | 30 | 24.98 | 24.9387 | 24.2 | 29.7 | 24.3 | 27.4 | 24.6 | 24.7 | 26.2 | 24.3 | 25.8 | 24.4 | 24.5 | 24.1467 | 26.039 | 24.6323 |  | 124
| 1 | 6 | 24.856 | 26 | 25.6 | 25.1 | 34.823 | 24 | 27.5 | 22.9 | 24.1 | 64 | 35.6 | 23 | 26 | 29 | 25.35 | 25.083 | 23.6 | 25.7 | 25.8 | 24.8 | 24.5 | 24.7 | 25.6 | 23.3 | 25.1 | 24.5 | 24.6 | 23.6 | 26.473 | 24.0922
| Mean |  | 24.4752 | 26.1333 | 25.5 | 24.15 | 35.3304 | 23.75 | 26.9167 | 23.6167 | 24.35 | 67 | 35.65 | 24 | 26.5 | 27.5 | 25.32 | 25.0667 | 23.7667 | 27.5167 | 25.3833 | 25.9833 | 25.25 | 25.45 | 25.2333 | 23.6 | 25.6 | 24.7167 | 24.6 | 24.21 | 26.4112 | 24.3305
| Median |  | 24.407 | 26.1 | 25.45 | 23.9 | 34.8757 | 23.75 | 26.75 | 23.9 | 24.35 | 68.5 | 35.75 | 24 | 26.5 | 28 | 25.345 | 25.066 | 23.7 | 27.45 | 25.7 | 25.9 | 25.2 | 25.6 | 25.25 | 23.6 | 25.45 | 24.65 | 24.6 | 24.0667 | 26.3975 | 24.2327
| Std Dev. |  | 0.2946 | 0.45019 | 0.42426 | 0.6442 | 2.1941 | 0.5244 | 1.0342 | 0.51929 | 0.80932 | 4.1952 | 1.8491 | 0.70711 | 0.54772 | 2.2583 | 0.18396 | 0.18828 | 0.2582 | 1.9763 | 0.61455 | 1.3333 | 0.6775 | 0.63797 | 1.1448 | 0.75366 | 0.61319 | 0.30605 | 0.22804 | 0.57053 | 0.28262 | 0.32201
| Rel.Std.Dev. |  | 1.204% | 1.723% | 1.664% | 2.668% | 6.21% | 2.208% | 3.842% | 2.199% | 3.324% | 6.262% | 5.187% | 2.946% | 2.067% | 8.212% | 0.7265% | 0.7511% | 1.086% | 7.182% | 2.421% | 5.131% | 2.683% | 2.507% | 4.537% | 3.193% | 2.395% | 1.238% | 0.927% | 2.357% | 1.07% | 1.323%
| PDM3 |  | -2.496% | 4.11% | 1.587% | -3.791% | 40.75% | -5.385% | 7.231% | -5.916% | -2.994% | 166.9% | 42.02% | -4.389% | 5.571% | 9.555% | 0.87% | -0.1391% | -5.318% | 9.621% | 1.122% | 3.513% | 0.5911% | 1.388% | 0.5247% | -5.982% | 1.985% | -1.534% | -1.998% | -3.552% | 5.217% | -3.072%
| Z-Score (Absolute) |  | 0.54 | 0.56 | 0.14 | 0.75 | 6.68 | 1.02 | 1.08 | 1.11 | 0.62 | 27.72 | 6.89 | 0.85 | 0.81 | 1.47 | 0.02 | 0.15 | 1.01 | 1.48 | 0.07 | 0.46 | 0.02 | 0.11 | 0.03 | 1.12 | 0.21 | 0.38 | 0.46 | 0.71 | 0.75 | 0.63
| Analytical results for Ta in OREAS 45f (Certified Value 1.66 ppm) | [32 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 1.75 | 1.49 | 1.75 | 1.9136 | 1.6 | 1.55 | 1.4 | 1.92 | < 5 | 1.48 | 1.1 | 1.4 | 1.4 | 1.54 | 2.0134 | 1.81 | 2.2 | 1.57 | 1.42 | 1.72 | 1.62 | 1.85 | 1.86 | 1.43 | 1.98 | 1.4667 | 6.319 | 1.6864 |  |  |  | 1
| 1 | 2 | 1.78 | 1.52 | 1.83 | 1.9516 | 1.7 | 1.71 | 1.4 | 2.47 | 6 | 1.4 | 1.2 | 1.5 | 1.5 | 1.54 | 2.1 | 1.88 | 2.2 | 1.54 | 1.45 | 1.69 | 1.61 | 1.82 | 2.02 | 1.48 | 1.69 | 1.5367 | 5.855 | 1.7444 |  |  |  | 22
| 1 | 3 | 1.65 | 1.49 | 1.76 | 1.9033 | 1.4 | 1.64 | 1.35 | 1.7 | 6 | 1.44 | 1.3 | 1.5 | 1.4 | 1.51 | 2.374 | 1.77 | 2.06 | 1.55 | 1.4 | 1.62 | 1.56 | 1.75 | 1.91 | 1.45 | 1.76 | 1.46 | 6.108 | 1.7035 |  |  |  | 16
| 1 | 4 | 1.77 | 1.56 | 1.78 | 1.9914 | 1.6 | 1.54 | 1.42 | 2 | 6 | 1.59 | 1.1 | 1.5 | 1.6 | 1.51 | 2.2872 | 1.76 | 2.23 | 1.48 | 1.54 | 1.66 | 1.58 | 1.76 | 2.01 | 1.47 | 1.89 | 1.52 | 5.893 | 1.7267 |  |  |  | 1.6552
| 1 | 5 | 1.61 | 1.64 | 1.75 | 1.9892 | 1.4 | 1.56 | 1.37 | 1.42 | 7 | 1.43 | 1.2 | 1.5 | 1.4 | 1.56 | 2.1491 | 1.91 | 2.2 | 1.46 | 1.46 | 1.66 | 1.53 | 1.75 | 1.84 | 1.44 | 1.92 | 1.4667 | 5.988 | 1.6902 |  |  |  | 125
| 1 | 6 | 1.63 | 1.54 | 1.8 | 1.9236 | 1.6 | 1.44 | 1.36 | 2 | 5 | 1.44 | 1.1 | 1.5 | 1.5 | 1.52 | 2.3067 | 1.87 | 2.18 | 1.54 | 1.66 | 1.67 | 1.58 | 1.86 | 1.89 | 1.45 | 1.83 | 1.4667 | 6.015 | 1.7115
| Mean |  | 1.6983 | 1.54 | 1.7783 | 1.9455 | 1.55 | 1.5733 | 1.3833 | 1.9183 | 6 | 1.4633 | 1.1667 | 1.4833 | 1.4667 | 1.53 | 2.2051 | 1.8333 | 2.1783 | 1.5233 | 1.4883 | 1.67 | 1.58 | 1.7983 | 1.9217 | 1.4533 | 1.845 | 1.4861 | 6.0297 | 1.7105
| Median |  | 1.7 | 1.53 | 1.77 | 1.9376 | 1.6 | 1.555 | 1.385 | 1.96 | 6 | 1.44 | 1.15 | 1.5 | 1.45 | 1.53 | 2.2182 | 1.84 | 2.2 | 1.54 | 1.455 | 1.665 | 1.58 | 1.79 | 1.9 | 1.45 | 1.86 | 1.4667 | 6.0015 | 1.7075
| Std Dev. |  | 0.076529 | 0.056214 | 0.031885 | 0.03829 | 0.12247 | 0.092448 | 0.027325 | 0.35034 | 0.70711 | 0.067132 | 0.08165 | 0.040825 | 0.08165 | 0.02 | 0.13894 | 0.062183 | 0.060139 | 0.043205 | 0.096833 | 0.033466 | 0.032863 | 0.051153 | 0.076267 | 0.018619 | 0.1071 | 0.033228 | 0.16792 | 0.022156
| Rel.Std.Dev. |  | 4.506% | 3.65% | 1.793% | 1.968% | 7.902% | 5.876% | 1.975% | 18.26% | 11.79% | 4.588% | 6.999% | 2.752% | 5.567% | 1.307% | 6.301% | 3.392% | 2.761% | 2.836% | 6.506% | 2.004% | 2.08% | 2.844% | 3.969% | 1.281% | 5.805% | 2.236% | 2.785% | 1.295%
| PDM3 |  | 2.609% | -6.957% | 7.442% | 17.54% | -6.353% | -4.943% | -16.42% | 15.9% | 262.5% | -11.59% | -29.51% | -10.38% | -11.39% | -7.561% | 33.23% | 10.77% | 31.61% | -7.964% | -10.08% | 0.897% | -4.541% | 8.651% | 16.1% | -12.19% | 11.47% | -10.21% | 264.3% | 3.341%
| Z-Score (Absolute) |  | 0.31 | 0.36 | 0.65 | 1.35 | 0.32 | 0.22 | 1.02 | 1.24 | 15.98 | 0.68 | 1.93 | 0.6 | 0.67 | 0.4 | 2.44 | 0.88 | 2.33 | 0.43 | 0.58 | 0.19 | 0.19 | 0.73 | 1.25 | 0.72 | 0.93 | 0.59 | 18.56 | 0.36
| Analytical results for Tb in OREAS 45f (Certified Value 0.37 ppm) | [32 empty] | Certified
Tb | Round | Replicate | 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab; 25: Lab |  |  |  | 1
| No | No | 04: 02; 05: 04; 07: 05; 10: 12; 12: 13; 13: 15; 14: 18; 15: 23; 17: 24; 18: 25; 19: 27; 20: 33; 21: 34; 23: 38; 24: 39; 25: 40 |  |  |  | ppm
|  |  | 04: 4A*MS; 05: 4A*MS; 07: 4A*OES/MS; 10: 4A*MS; 12: 4A*OES/MS; 13: 4A*MS; 14: 4A*MS; 15: 4A*MS; 17: 4A*OES/MS; 18: 4A*MS; 19: 4A*OES/MS; 20: 4A*MS; 21: 4A*MS; 23: 4A*MS; 24: 5A*OES; 25: 4A*MS |  |  |  | 2
3
| 1 | 1 | 04: 0.3922; 05: 0.34; 07: 0.41; 10: 0.36; 12: 0.37; 13: 0.4; 14: 0.48; 15: 0.34; 17: 0.37327; 18: 0.298; 19: 0.37; 20: 0.4; 21: 0.29; 23: 0.34333; 24: 0.3; 25: 0.4004 |  |  |  | 1
| 1 | 2 | 04: 0.3926; 05: 0.36; 07: 0.41; 10: 0.36; 12: 0.35; 13: 0.36; 14: 0.45; 15: 0.35; 17: 0.37302; 18: 0.314; 19: 0.38; 20: 0.35; 21: 0.3; 23: 0.35667; 24: 0.6; 25: 0.3948 |  |  |  | #N/A
| 1 | 3 | 04: 0.3863; 05: 0.33; 07: 0.39; 10: 0.36; 12: 0.36; 13: 0.4; 14: 0.48; 15: 0.35; 17: 0.35542; 18: 0.304; 19: 0.36; 20: 0.39; 21: 0.28; 23: 0.33667; 24: 0.5; 25: 0.3961 |  |  |  | 16
| 1 | 4 | 04: 0.3894; 05: 0.36; 07: 0.4; 10: 0.36; 12: 0.36; 13: 0.37; 14: 0.5; 15: 0.35; 17: 0.35031; 18: 0.31; 19: 0.36; 20: 0.43; 21: 0.27; 23: 0.35; 24: 0.3; 25: 0.3753 |  |  |  | 0.3716
| 1 | 5 | 04: 0.3861; 05: 0.34; 07: 0.42; 10: 0.38; 12: 0.38; 13: 0.33; 14: 0.44; 15: 0.35; 17: 0.37355; 18: 0.315; 19: 0.36; 20: 0.45; 21: 0.3; 23: 0.34667; 24: 0.6; 25: 0.4112 |  |  |  | 126
| 1 | 6 | 04: 0.3806; 05: 0.36; 07: 0.38; 10: 0.36; 12: 0.35; 13: 0.38; 14: 0.43; 15: 0.36; 17: 0.38246; 18: 0.306; 19: 0.35; 20: 0.43; 21: 0.3; 23: 0.34667; 24: 0.4; 25: 0.392
| Mean |  | 04: 0.38787; 05: 0.34833; 07: 0.40167; 10: 0.36333; 12: 0.36167; 13: 0.37333; 14: 0.46333; 15: 0.35; 17: 0.36801; 18: 0.30783; 19: 0.36333; 20: 0.40833; 21: 0.29; 23: 0.34667; 24: 0.45; 25: 0.39497
| Median |  | 04: 0.38785; 05: 0.35; 07: 0.405; 10: 0.36; 12: 0.36; 13: 0.375; 14: 0.465; 15: 0.35; 17: 0.37314; 18: 0.308; 19: 0.36; 20: 0.415; 21: 0.295; 23: 0.34667; 24: 0.45; 25: 0.39545
| Std Dev. |  | 04: 0.0045147; 05: 0.013292; 07: 0.01472; 10: 0.008165; 12: 0.01169; 13: 0.026583; 14: 0.027325; 15: 0.0063246; 17: 0.01236; 18: 0.0064627; 19: 0.010328; 20: 0.036009; 21: 0.012649; 23: 0.0066667; 24: 0.13784; 25: 0.011747
| Rel.Std.Dev. |  | 04: 1.164%; 05: 3.816%; 07: 3.665%; 10: 2.247%; 12: 3.232%; 13: 7.121%; 14: 5.898%; 15: 1.807%; 17: 3.359%; 18: 2.099%; 19: 2.843%; 20: 8.819%; 21: 4.362%; 23: 1.923%; 24: 30.63%; 25: 2.974%
| PDM3 |  | 04: 4.378%; 05: -6.261%; 07: 8.092%; 10: -2.224%; 12: -2.673%; 13: 0.4671%; 14: 24.69%; 15: -5.812%; 17: -0.9666%; 18: -17.16%; 19: -2.224%; 20: 9.886%; 21: -21.96%; 23: -6.709%; 24: 21.1%; 25: 6.289%
| Z-Score (Absolute) |  | 04: 0.99; 05: 0.61; 07: 1.55; 10: 0; 12: 0.07; 13: 0.4; 14: 4.05; 15: 0.54; 17: 0.19; 18: 2.25; 19: 0; 20: 1.82; 21: 2.97; 23: 0.67; 24: NA; 25: 1.28
Results from laboratory 39 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 45f (Indicative Value 0.08 ppm) | [32 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | [6 empty] | 3
3
| 1 | 1 | 0.1 | < 0.05 | 0.1 | 3.5041 | < 0.2 | < 0.01 | 0.05 | 0.17 | 0.1 | < 0.2 | < 0.5 | < 0.5 | 0.065768 | < 0.05 | 0.27 | 0.06 | 0.06 | < 0.05 | 0.09 | < 0.1 | 0.08 | 0.08 | 0.11 | < 0.1 | 0.058 | [6 empty] | 1
| 1 | 2 | 0.09 | 0.06 | 0.1 | 3.7331 | < 0.2 | 0.11 | 0.06 | 0.13 | 0.1 | < 0.2 | < 0.5 | < 0.5 | 0.055878 | < 0.05 | 0.19 | 0.04 | 0.06 | < 0.05 | 0.15 | < 0.1 | 0.06 | 0.11 | 0.14 | < 0.1 | 0.072 | [6 empty] | 9
| 1 | 3 | 0.08 | 0.07 | 0.1 | 3.6508 | < 0.2 | 0.07 | 0.07 | 0.13 | < 0.1 | < 0.2 | < 0.5 | < 0.5 | 0.07635 | 0.126 | 0.16 | 0.06 | 0.11 | 0.13 | 0.08 | < 0.1 | 0.07 | 0.09 | 0.1 | < 0.1 | 0.054 | [6 empty] | 16
| 1 | 4 | 0.08 | 0.06 | 0.1 | 3.737 | < 0.2 | 0.21 | 0.07 | 0.3 | 0.1 | < 0.2 | < 0.5 | < 0.5 | 0.09349 | < 0.05 | 0.22 | 0.07 | 0.06 | 0.07 | 0.08 | < 0.1 | 0.07 | 0.1 | 0.09 | < 0.1 | 0.061 | [6 empty] | 0.080383
| 1 | 5 | 0.1 | 0.06 | < 0.1 | 3.6716 | < 0.2 | < 0.01 | 0.07 | 0.26 | < 0.1 | < 0.2 | < 0.5 | < 0.5 | 0.07639 | 0.114 | 0.7 | 0.08 | 0.07 | 0.05 | 0.1 | < 0.1 | 0.06 | 0.08 | 0.08 | < 0.1 | 0.062 | [6 empty] | 23
| 1 | 6 | 0.09 | 0.08 | 0.1 | 3.3123 | < 0.2 | 0.18 | 0.06 | 0.17 | < 0.1 | < 0.2 | < 0.5 | < 0.5 | 0.064259 | 0.052 | 0.14 | 0.07 | 0.08 | 0.12 | 0.09 | < 0.1 | 0.11 | 0.13 | 0.1 | < 0.1 | 0.054
| Mean |  | 0.09 | 0.066 | 0.1 | 3.6015 |  | 0.1425 | 0.063333 | 0.19333 | 0.1 |  |  |  | 0.072023 | 0.097333 | 0.28 | 0.063333 | 0.073333 | 0.0925 | 0.098333 |  | 0.075 | 0.098333 | 0.10333 |  | 0.060167
| Median |  | 0.09 | 0.06 | 0.1 | 3.6612 |  | 0.145 | 0.065 | 0.17 | 0.1 |  |  |  | 0.071059 | 0.114 | 0.205 | 0.065 | 0.065 | 0.095 | 0.09 |  | 0.07 | 0.095 | 0.1 |  | 0.0595
| Std Dev. |  | 0.0089443 | 0.0089443 | 0 | 0.16502 |  | 0.063966 | 0.008165 | 0.070616 | 1.6997e-17 |  |  |  | 0.013105 | 0.039716 | 0.21081 | 0.013663 | 0.019664 | 0.038622 | 0.026394 |  | 0.018708 | 0.019408 | 0.020656 |  | 0.0067057
| Rel.Std.Dev. |  | 9.938% | 13.55% | 0% | 4.582% |  | 44.89% | 12.89% | 36.53% | 1.7e-14% |  |  |  | 18.2% | 40.8% | 75.29% | 21.57% | 26.81% | 41.75% | 26.84% |  | 24.94% | 19.74% | 19.99% |  | 11.15%
| PDM3 |  | 11.96% | -17.89% | 24.41% | 4380% |  | 77.28% | -21.21% | 140.5% | 24.41% |  |  |  | -10.4% | 21.09% | 248.3% | -21.21% | -8.77% | 15.07% | 22.33% |  | -6.696% | 22.33% | 28.55% |  | -25.15%
| Z-Score (Absolute) |  | 0 | 0.78 | NA | 88.79 | 0.25 | 0.17 | 0.67 | 2.61 | NA | 0.25 | 4.05 | 4.05 | 0.45 | 0.73 | 4.8 | 0.67 | 0.42 | 0.51 | 0.21 | 1.01 | 0.38 | 0.21 | 0.34 | 1.01 | 0.75
Results from laboratories 7 & 15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 45f (Certified Value 9.99 ppm) | [32 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 11.353 | 9.6 | 10.2 | 10.4 | 12.1412 | 9.3 | 10.2 | 8.44 | 8.26 | 11.3 | 10.9 | 9 | 9.64 | 10.0099 | 10.8 | 9.8 | 9.63 | 9.37 | 11.05 | 9.75 | 11.5 | 7.2 | 8.98 | 11.2 | 9.9367 | 9.51 | 9.4933 |  |  |  |  | 1
| 1 | 2 | 11.21 | 10.1 | 10.3 | 10.6 | 11.2723 | 9.4 | 10.1 | 9 | 8.26 | 10.4 | 10.4 | 9.6 | 9.72 | 10.0163 | 10.7 | 9.8 | 10.15 | 9.75 | 9.76 | 10.1 | 11.3 | 7.8 | 9.25 | 11.2 | 10.5467 | 8.923 | 9.8163 |  |  |  |  | #N/A
| 1 | 3 | 11.26 | 9.5 | 10.1 | 10.5 | 11.0857 | 9.2 | 10.6 | 8.3 | 9.67 | 11.4 | 10.3 | 8.5 | 9.54 | 9.6988 | 10.6 | 9.6 | 10.25 | 9.5 | 9.94 | 9.74 | 10.8 | 7 | 9.25 | 11 | 9.8967 | 9.561 | 9.0895 |  |  |  |  | 16
| 1 | 4 | 11.424 | 10.2 | 10.4 | 10.2 | 12.1522 | 9.3 | 9.4 | 9.01 | 9.05 | 14.6 | 10.5 | 9.8 | 9.76 | 9.8109 | 10.4 | 9.8 | 9.04 | 9.66 | 10.2 | 9.65 | 10.8 | 6.9 | 9.69 | 10.7 | 10.0367 | 8.329 | 9.1629 |  |  |  |  | 9.9911
| 1 | 5 | 11.294 | 10.3 | 10.7 | 10.2 | 11.8886 | 9 | 10.2 | 8.06 | 7.8 | 11.1 | 11.6 | 9.3 | 9.76 | 9.8057 | 10.9 | 9.4 | 8.58 | 9.6 | 10.25 | 8.79 | 10.9 | 7.7 | 9.7 | 10.9 | 9.7533 | 9.162 | 9.3342 |  |  |  |  | 127
| 1 | 6 | 11.63 | 10.1 | 9.8 | 10.5 | 11.6599 | 9.1 | 10.7 | 7.38 | 9 | 12 | 10.6 | 9.4 | 9.87 | 9.9339 | 10.8 | 9.4 | 9.2 | 9.66 | 10.75 | 8.87 | 13.5 | 7.7 | 9.34 | 10.8 | 9.9033 | 9.367 | 9.1819
| Mean |  | 11.3618 | 9.9667 | 10.25 | 10.4 | 11.7 | 9.2167 | 10.2 | 8.365 | 8.6733 | 11.8 | 10.7167 | 9.2667 | 9.715 | 9.8792 | 10.7 | 9.6333 | 9.475 | 9.59 | 10.325 | 9.4833 | 11.4667 | 7.3833 | 9.3683 | 10.9667 | 10.0122 | 9.142 | 9.3463
| Median |  | 11.3235 | 10.1 | 10.25 | 10.45 | 11.7742 | 9.25 | 10.2 | 8.37 | 8.63 | 11.35 | 10.55 | 9.35 | 9.74 | 9.8724 | 10.75 | 9.7 | 9.415 | 9.63 | 10.225 | 9.695 | 11.1 | 7.45 | 9.295 | 10.95 | 9.92 | 9.2645 | 9.2581
| Std Dev. |  | 0.15092 | 0.33267 | 0.30166 | 0.16733 | 0.44646 | 0.1472 | 0.46043 | 0.6152 | 0.68503 | 1.4656 | 0.47924 | 0.46332 | 0.11345 | 0.12767 | 0.17889 | 0.19664 | 0.65516 | 0.13565 | 0.48887 | 0.52952 | 1.0367 | 0.39707 | 0.28039 | 0.20656 | 0.27719 | 0.4628 | 0.27176
| Rel.Std.Dev. |  | 1.328% | 3.338% | 2.943% | 1.609% | 3.816% | 1.597% | 4.514% | 7.354% | 7.898% | 12.42% | 4.472% | 5% | 1.168% | 1.292% | 1.672% | 2.041% | 6.915% | 1.414% | 4.735% | 5.584% | 9.041% | 5.378% | 2.993% | 1.884% | 2.769% | 5.062% | 2.908%
| PDM3 |  | 13.72% | -0.245% | 2.591% | 4.092% | 17.1% | -7.752% | 2.09% | -16.28% | -13.19% | 18.1% | 7.262% | -7.251% | -2.764% | -1.12% | 7.095% | -3.581% | -5.166% | -4.015% | 3.342% | -5.083% | 14.77% | -26.1% | -6.234% | 9.764% | 0.211% | -8.499% | -6.454%
| Z-Score (Absolute) |  | 1.88 | 0.11 | 0.47 | 0.66 | 2.3 | 0.84 | 0.41 | 1.92 | 1.53 | 2.43 | 1.06 | 0.78 | 0.21 | 0 | 1.04 | 0.31 | 0.51 | 0.37 | 0.56 | 0.5 | 2.01 | 3.16 | 0.65 | 1.38 | 0.17 | 0.93 | 0.67
| Analytical results for Ti in OREAS 45f (Certified Value 1.08 wt.%) | [32 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  |  |  | 2
3
| 1 | 1 | 1.0562 | 1.09 | 1.08 | 1.15 | 0.91575 | 1.13 | 1.27 | 1.03 | 1.03 | 0.97 | 1.07 | 1.041 | 1.1196 | 1.0848 | 0.914 | 1.1 | 1.035 | 1.06 | 1.095 | 1.11 | 1.01 | 1.12 | 1.06 | 1.11 | 1.08 | 1.0357 | 0.81828 | 1.0872 |  |  |  | 1
| 1 | 2 | 1.0664 | 1.09 | 1.11 | 1.17 | 0.9033 | 1.14 | 1.13 | 1.03 | 1.04 | 0.99 | 1.044 | 1.083 | 1.0983 | 1.1128 | 0.928 | 1.06 | 1.075 | 1.06 | 1.17 | 1.06 | 1.03 | 1.08 | 1.09 | 1.06 | 1.07 | 1.0155 | 0.82142 | 1.1004 |  |  |  | #N/A
| 1 | 3 | 1.0516 | 1.06 | 1.1 | 1.15 | 0.8891 | 1.13 | 1.13 | 0.993 | 1.01 | 1.01 | 1.08 | 1.041 | 1.0854 | 1.1049 | 0.899 | 1.08 | 1.075 | 1.085 | 1.08 | 1.1 | 1.06 | 1.02 | 1.065 | 1.08 | 1.05 | 1.0284 | 0.81997 | 1.109 |  |  |  | 16
| 1 | 4 | 1.0555 | 1.04 | 1.12 | 1.14 | 0.89035 | 1.15 | 1.14 | 1.04 | 1.02 | 1.07 | 1.08 | 1.085 | 1.1001 | 1.101 | 0.889 | 1.1 | 1.045 | 1.065 | 1.13 | 1.075 | 1.09 | 1.09 | 1.055 | 1.08 | 1.07 | 1.0204 | 0.8075 | 1.1011 |  |  |  | 1.0765
| 1 | 5 | 1.0744 | 1.06 | 1.17 | 1.14 | 0.9036 | 1.14 | 1.12 | 1.025 | 1.03 | 1.09 | 1.079 | 1.031 | 1.1155 | 1.0992 | 0.921 | 1.07 | 1.05 | 1.085 | 1.1 | 1.03 | 1.07 | 1.07 | 1.05 | 1.07 | 1.04 | 1.0475 | 0.80686 | 1.0847 |  |  |  | 128
| 1 | 6 | 1.0562 | 1.08 | 1.12 | 1.14 | 0.90955 | 1.15 | 1.07 | 1.01 | 1.03 | 0.79 | 1.074 | 1.08 | 1.0917 | 1.1248 | 0.9 | 1.02 | 1.045 | 1.08 | 1.145 | 1.045 | 1.03 | 1.09 | 1.035 | 1.08 | 1.06 | 1.0266 | 0.8019 | 1.1
| Mean |  | 1.06 | 1.07 | 1.1167 | 1.1483 | 0.90194 | 1.14 | 1.1433 | 1.0213 | 1.0267 | 0.98667 | 1.0712 | 1.0602 | 1.1018 | 1.1046 | 0.9085 | 1.0717 | 1.0542 | 1.0725 | 1.12 | 1.07 | 1.0483 | 1.0783 | 1.0592 | 1.08 | 1.0617 | 1.029 | 0.81265 | 1.0971
| Median |  | 1.0562 | 1.07 | 1.115 | 1.145 | 0.90345 | 1.14 | 1.13 | 1.0275 | 1.03 | 1 | 1.0765 | 1.0605 | 1.0992 | 1.103 | 0.907 | 1.075 | 1.0475 | 1.0725 | 1.115 | 1.0675 | 1.045 | 1.085 | 1.0575 | 1.08 | 1.065 | 1.0275 | 0.81289 | 1.1002
| Std Dev. |  | 0.008579 | 0.02 | 0.030111 | 0.01169 | 0.010511 | 0.0089443 | 0.066833 | 0.01699 | 0.010328 | 0.1069 | 0.013891 | 0.024967 | 0.013346 | 0.013489 | 0.014896 | 0.029944 | 0.016857 | 0.012145 | 0.034205 | 0.031145 | 0.029944 | 0.033116 | 0.01828 | 0.016733 | 0.01472 | 0.011413 | 0.0082207 | 0.0092646
| Rel.Std.Dev. |  | 0.8093% | 1.869% | 2.696% | 1.018% | 1.165% | 0.7846% | 5.845% | 1.664% | 1.006% | 10.83% | 1.297% | 2.355% | 1.211% | 1.221% | 1.64% | 2.794% | 1.599% | 1.132% | 3.054% | 2.911% | 2.856% | 3.071% | 1.726% | 1.549% | 1.386% | 1.109% | 1.012% | 0.8445%
| PDM3 |  | -1.527% | -0.6024% | 3.733% | 6.674% | -16.21% | 5.9% | 6.21% | -5.123% | -4.628% | -8.344% | -0.494% | -1.516% | 2.349% | 2.611% | -15.6% | -0.4476% | -2.073% | -0.3702% | 4.042% | -0.6024% | -2.615% | 0.1717% | -1.609% | 0.3265% | -1.377% | -4.408% | -24.51% | 1.913%
| Z-Score (Absolute) |  | 0.23 | 0 | 1.07 | 1.8 | 3.85 | 1.6 | 1.68 | 1.12 | 0.99 | 1.91 | 0.03 | 0.23 | 0.73 | 0.79 | 3.7 | 0.04 | 0.36 | 0.06 | 1.15 | 0 | 0.5 | 0.19 | 0.25 | 0.23 | 0.19 | 0.94 | 5.9 | 0.62
| Analytical results for Tl in OREAS 45f (Certified Value 0.2 ppm) | [32 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | < 0.25 | 0.19 | 0.21 | 0.18 | 0.29717 | 0.2 | 0.22 | 0.189 | 0.18 | 0.2 | 0.1 | < 0.5 | < 0.5 | 0.19 | 0.21069 | 0.224 | 0.2 | 0.204 | 0.201 | 0.23 | 0.186 | 0.2 | 0.21 | 0.202 | 0.14 | 0.18667 | 0.212 | 0.2079 |  |  |  | 1
| 1 | 2 | < 0.25 | 0.21 | 0.22 | 0.21 | 0.27796 | 0.2 | 0.21 | 0.183 | 0.17 | < 0.2 | 0.1 | < 0.5 | < 0.5 | 0.2 | 0.20797 | 0.227 | 0.2 | 0.211 | 0.197 | 0.23 | 0.181 | 0.2 | 0.22 | 0.22 | 0.15 | 0.19333 | 0.181 | 0.2008 |  |  |  | 10
| 1 | 3 | < 0.25 | 0.18 | 0.2 | 0.18 | 0.26762 | 0.2 | 0.21 | 0.18 | 0.21 | < 0.2 | 0.1 | < 0.5 | < 0.5 | 0.19 | 0.20632 | 0.251 | 0.2 | 0.201 | 0.196 | 0.24 | 0.183 | 0.2 | 0.21 | 0.213 | 0.14 | 0.18667 | 0.216 | 0.2 |  |  |  | 16
| 1 | 4 | 0.25227 | 0.2 | 0.21 | 0.2 | 0.31728 | 0.2 | 0.2 | 0.177 | 0.19 | 0.2 | 0.1 | < 0.5 | < 0.5 | 0.18 | 0.20679 | 0.223 | 0.2 | 0.205 | 0.194 | 0.22 | 0.184 | 0.2 | 0.22 | 0.227 | 0.13 | 0.18667 | 0.175 | 0.1953 |  |  |  | 0.2017
| 1 | 5 | < 0.25 | 0.2 | 0.21 | 0.19 | 0.27208 | 0.2 | 0.22 | 0.172 | 0.19 | < 0.2 | 0.1 | < 0.5 | < 0.5 | 0.19 | 0.21128 | 0.234 | 0.2 | 0.195 | 0.193 | 0.22 | 0.175 | 0.2 | 0.22 | 0.215 | 0.13 | 0.18333 | 0.182 | 0.1976 |  |  |  | 129
| 1 | 6 | < 0.25 | 0.2 | 0.21 | 0.2 | 0.31912 | 0.2 | 0.19 | 0.188 | 0.19 | < 0.2 | < 0.1 | < 0.5 | < 0.5 | 0.18 | 0.20634 | 0.223 | 0.2 | 0.21 | 0.2 | 0.24 | 0.172 | 0.2 | 0.24 | 0.217 | 0.13 | 0.18667 | 0.213 | 0.1999
| Mean |  | 0.25227 | 0.19667 | 0.21 | 0.19333 | 0.29187 | 0.2 | 0.20833 | 0.1815 | 0.18833 | 0.2 | 0.1 |  |  | 0.18833 | 0.20823 | 0.23033 | 0.2 | 0.20433 | 0.19683 | 0.23 | 0.18017 | 0.2 | 0.22 | 0.21567 | 0.13667 | 0.18722 | 0.1965 | 0.20025
| Median |  | 0.25227 | 0.2 | 0.21 | 0.195 | 0.28757 | 0.2 | 0.21 | 0.1815 | 0.19 | 0.2 | 0.1 |  |  | 0.19 | 0.20738 | 0.2255 | 0.2 | 0.2045 | 0.1965 | 0.23 | 0.182 | 0.2 | 0.22 | 0.216 | 0.135 | 0.18667 | 0.197 | 0.19995
| Std Dev. |  |  | 0.010328 | 0.0063246 | 0.012111 | 0.022757 | 3.0405e-17 | 0.01169 | 0.0065345 | 0.013292 | 0 | 0 |  |  | 0.0075277 | 0.0022243 | 0.010948 | 3.0405e-17 | 0.0059217 | 0.0031885 | 0.0089443 | 0.0054924 | 3.0405e-17 | 0.010954 | 0.0082865 | 0.008165 | 0.0032773 | 0.019003 | 0.0042552
| Rel.Std.Dev. |  |  | 5.252% | 3.012% | 6.264% | 7.797% | 1.52e-14% | 5.611% | 3.6% | 7.057% | 0% | 0% |  |  | 3.997% | 1.068% | 4.753% | 1.52e-14% | 2.898% | 1.62% | 3.889% | 3.049% | 1.52e-14% | 4.979% | 3.842% | 5.974% | 1.75% | 9.671% | 2.125%
| PDM3 |  | 25.07% | -2.498% | 4.113% | -4.15% | 44.7% | -0.8451% | 3.286% | -10.02% | -6.629% | -0.8451% | -50.42% |  |  | -6.629% | 3.236% | 14.19% | -0.8451% | 1.303% | -2.415% | 14.03% | -10.68% | -0.8451% | 9.07% | 6.922% | -32.24% | -7.18% | -2.58% | -0.7211%
| Z-Score (Absolute) |  | 2.8 | 0.19 | 0.5 | 0.36 | 4.74 | NA | 0.42 | 0.97 | 0.62 | NA | NA | 2.58 | 2.58 | 0.62 | 0.41 | 1.56 | NA | 0.21 | 0.18 | 1.54 | 1.04 | NA | 1.02 | 0.8 | 3.29 | 0.67 | 0.19 | 0
Results from laboratories 12, 18, 19, 27 & 33 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 45f (Certified Value 0.2 ppm) | [32 empty] | Certified
Tm | Round | Replicate | 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab |  |  |  | 1
| No | No | 04: 02; 05: 12; 07: 13; 10: 15; 12: 18; 13: 23; 14: 24; 15: 27; 17: 33; 18: 38; 19: 39; 20: 40 |  |  |  | ppm
|  |  | 04: 4A*MS; 05: 4A*MS; 07: 4A*OES/MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*OES/MS; 17: 4A*MS; 18: 4A*MS; 19: 5A*MS; 20: 4A*MS |  |  |  | 2
2
| 1 | 1 | 04: < 0.25; 05: 0.2; 07: 0.21; 10: 0.22; 12: 0.2; 13: 0.18; 14: 0.17368; 15: 0.19; 17: 0.24; 18: 0.18; 19: 0.158; 20: 0.2429 |  |  |  | 1
| 1 | 2 | 04: < 0.25; 05: 0.22; 07: 0.24; 10: 0.23; 12: 0.19; 13: 0.18; 14: 0.19403; 15: 0.2; 17: 0.22; 18: 0.19; 19: 0.142; 20: 0.2468 |  |  |  | #N/A
| 1 | 3 | 04: < 0.25; 05: 0.18; 07: 0.2; 10: 0.28; 12: 0.2; 13: 0.17; 14: 0.18699; 15: 0.19; 17: 0.23; 18: 0.19667; 19: 0.161; 20: 0.2496 |  |  |  | 16
| 1 | 4 | 04: < 0.25; 05: 0.22; 07: 0.2; 10: 0.24; 12: 0.23; 13: 0.17; 14: 0.19336; 15: 0.19; 17: 0.25; 18: 0.18667; 19: 0.141; 20: 0.2374 |  |  |  | 0.2024
| 1 | 5 | 04: < 0.25; 05: 0.2; 07: 0.19; 10: 0.23; 12: 0.2; 13: 0.19; 14: 0.17786; 15: 0.2; 17: 0.24; 18: 0.18667; 19: 0.158; 20: 0.2437 |  |  |  | 130
| 1 | 6 | 04: < 0.25; 05: 0.18; 07: 0.23; 10: 0.26; 12: 0.19; 13: 0.19; 14: 0.19052; 15: 0.19; 17: 0.23; 18: 0.18667; 19: 0.149; 20: 0.2299
| Mean |  | 05: 0.2; 07: 0.21167; 10: 0.24333; 12: 0.20167; 13: 0.18; 14: 0.18607; 15: 0.19333; 17: 0.235; 18: 0.18778; 19: 0.1515; 20: 0.24172
| Median |  | 05: 0.2; 07: 0.205; 10: 0.235; 12: 0.2; 13: 0.18; 14: 0.18876; 15: 0.19; 17: 0.235; 18: 0.18667; 19: 0.1535; 20: 0.2433
| Std Dev. |  | 05: 0.017889; 07: 0.019408; 10: 0.022509; 12: 0.01472; 13: 0.0089443; 14: 0.0084607; 15: 0.005164; 17: 0.010488; 18: 0.0054433; 19: 0.008735; 20: 0.0070935
| Rel.Std.Dev. |  | 05: 8.944%; 07: 9.169%; 10: 9.25%; 12: 7.299%; 13: 4.969%; 14: 4.547%; 15: 2.671%; 17: 4.463%; 18: 2.899%; 19: 5.766%; 20: 2.935%
| PDM3 |  | 05: -1.186%; 07: 4.578%; 10: 20.22%; 12: -0.3624%; 13: -11.07%; 14: -8.066%; 15: -4.48%; 17: 16.11%; 18: -7.225%; 19: -25.15%; 20: 19.43%
| Z-Score (Absolute) |  | 04: 3.05; 05: 0.14; 07: 0.64; 10: 1.99; 12: 0.21; 13: 0.71; 14: 0.45; 15: 0.14; 17: 1.63; 18: 0.38; 19: 1.92; 20: 1.92
| Analytical results for U in OREAS 45f (Certified Value 2.09 ppm) | [32 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 5A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.2493 | 1.8 | 2.17 | 1.85 | 2.3773 | 2.2 | 2.12 | 1.93 | 2.2 | 2 | 2 | 2.2 | 1.9 | 1.96 | 2.1271 | 2.21 | 2.13 | 1.98 | 2 | 2.1 | 2.17 | 2.02 | 2.2 | 1.9 | 2.45 | 1.9533 | 1.846 | 2.4301 |  |  |  | 1
| 1 | 2 | 2.2477 | 2 | 2.13 | 2.01 | 2.0403 | 2.2 | 2.24 | 1.88 | 2.19 | 1.85 | 2 | 2.3 | 2 | 1.91 | 2.1223 | 2.23 | 1.99 | 2.14 | 2.08 | 2.1 | 2.14 | 2.14 | 2.2 | 1.99 | 2.35 | 2.07 | 1.662 | 2.4456 |  |  |  | #N/A
| 1 | 3 | 2.2205 | 1.8 | 2.14 | 1.84 | 1.9814 | 2 | 2.18 | 1.86 | 2.6 | 1.91 | 2 | 2.2 | 1.8 | 1.92 | 2.2278 | 2.13 | 2.07 | 2.14 | 2.04 | 2 | 2.1 | 2 | 2.1 | 2.02 | 2.3 | 1.96 | 1.796 | 2.3173 |  |  |  | 16
| 1 | 4 | 2.2345 | 2 | 2.21 | 1.85 | 2.3736 | 2.3 | 2.07 | 1.91 | 2.4 | 2.03 | 2 | 2.2 | 2 | 1.91 | 2.2137 | 2.15 | 2.18 | 2.11 | 1.98 | 2.2 | 2.12 | 2.02 | 2.2 | 2.1 | 2.26 | 1.9567 | 1.561 | 2.3449 |  |  |  | 2.0888
| 1 | 5 | 2.2248 | 1.9 | 2.29 | 1.85 | 2.1335 | 2.2 | 2.08 | 1.87 | 2.3 | 1.84 | 2.2 | 2.3 | 1.9 | 1.95 | 2.1573 | 2.2 | 1.95 | 1.91 | 2.05 | 2.1 | 2.03 | 1.92 | 2.2 | 2.06 | 2.29 | 1.9367 | 1.728 | 2.3094 |  |  |  | 131
| 1 | 6 | 2.2319 | 1.9 | 2.15 | 1.99 | 2.1056 | 2.2 | 2.24 | 1.86 | 2.3 | 1.92 | 2.1 | 2.2 | 1.8 | 1.94 | 2.1726 | 2.2 | 1.9 | 2.01 | 2.02 | 2.1 | 2.06 | 2.06 | 2.3 | 1.97 | 2.26 | 1.95 | 1.741 | 2.3466
| Mean |  | 2.2348 | 1.9 | 2.1817 | 1.8983 | 2.1686 | 2.1833 | 2.155 | 1.885 | 2.3317 | 1.925 | 2.05 | 2.2333 | 1.9 | 1.9317 | 2.1701 | 2.1867 | 2.0367 | 2.0483 | 2.0283 | 2.1 | 2.1033 | 2.0267 | 2.2 | 2.0067 | 2.3183 | 1.9711 | 1.7223 | 2.3657
| Median |  | 2.2332 | 1.9 | 2.16 | 1.85 | 2.1195 | 2.2 | 2.15 | 1.875 | 2.3 | 1.915 | 2 | 2.2 | 1.9 | 1.93 | 2.165 | 2.2 | 2.03 | 2.06 | 2.03 | 2.1 | 2.11 | 2.02 | 2.2 | 2.005 | 2.295 | 1.955 | 1.7345 | 2.3457
| Std Dev. |  | 0.011744 | 0.089443 | 0.060139 | 0.079099 | 0.1687 | 0.098319 | 0.076354 | 0.02881 | 0.15237 | 0.077136 | 0.083666 | 0.05164 | 0.089443 | 0.02137 | 0.043666 | 0.038297 | 0.10838 | 0.095795 | 0.036009 | 0.063246 | 0.05164 | 0.072296 | 0.063246 | 0.070333 | 0.072503 | 0.049111 | 0.10077 | 0.058036
| Rel.Std.Dev. |  | 0.5255% | 4.708% | 2.757% | 4.167% | 7.779% | 4.503% | 3.543% | 1.528% | 6.535% | 4.007% | 4.081% | 2.312% | 4.708% | 1.106% | 2.012% | 1.751% | 5.322% | 4.677% | 1.775% | 3.012% | 2.455% | 3.567% | 2.875% | 3.505% | 3.127% | 2.492% | 5.851% | 2.453%
| PDM3 |  | 6.991% | -9.037% | 4.448% | -9.117% | 3.824% | 4.528% | 3.171% | -9.755% | 11.63% | -7.84% | -1.856% | 6.922% | -9.037% | -7.521% | 3.895% | 4.687% | -2.494% | -1.935% | -2.893% | 0.5381% | 0.6977% | -2.973% | 5.326% | -3.93% | 10.99% | -5.632% | -17.54% | 13.26%
| Z-Score (Absolute) |  | 0.98 | 1.07 | 0.65 | 1.08 | 0.57 | 0.66 | 0.49 | 1.16 | 1.57 | 0.92 | 0.15 | 0.97 | 1.07 | 0.88 | 0.58 | 0.68 | 0.23 | 0.16 | 0.29 | 0.15 | 0.17 | 0.3 | 0.77 | 0.42 | 1.49 | 0.64 | 2.16 | 1.78
| Analytical results for V in OREAS 45f (Certified Value 253 ppm) | [32 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*OES |  | 0
0
| 1 | 1 | 256.35 | 255 | 280 | 257 | 256.1269 | 265 | 262 | 245 | 186 | 256 | 252 | 245 | 271 | 244 | 246 | 251.1503 | 251 | 305 | 247 | 249 | 273 | 261 | 243 | 245 | 262 | 262 | 248 | 246.6667 | 244.478 | 249.1033 |  | 1
| 1 | 2 | 255.19 | 260 | 270 | 255 | 258.5139 | 265 | 260 | 242 | 194 | 258 | 241 | 240 | 256 | 247 | 253 | 256.7926 | 257 | 287 | 260 | 244 | 283 | 251 | 232 | 246 | 270 | 258 | 243 | 237.6667 | 242.208 | 253.0836 |  | #N/A
| 1 | 3 | 255.97 | 245 | 272 | 256 | 245.9475 | 270 | 256 | 239 | 191 | 256 | 258 | 255 | 261 | 244 | 257 | 251.657 | 253 | 295 | 259 | 250 | 267 | 262 | 250 | 235 | 263 | 261 | 240 | 238 | 237.579 | 254.449 |  | 16
| 1 | 4 | 253.13 | 253 | 282 | 250 | 253.3832 | 275 | 257 | 242 | 193 | 255 | 264 | 255 | 273 | 258 | 262 | 257.9404 | 256 | 302 | 254 | 249 | 275 | 256 | 229 | 246 | 261 | 255 | 248 | 235.6667 | 240.858 | 252.3027 |  | 252.5181
| 1 | 5 | 256.72 | 254 | 277 | 253 | 251.3111 | 265 | 259 | 245 | 183 | 257 | 243 | 260 | 268 | 244 | 259 | 256.5442 | 256 | 287 | 250 | 252 | 270 | 245 | 245 | 244 | 260 | 267 | 246 | 235 | 240.814 | 248.6 |  | 132
| 1 | 6 | 255.85 | 257 | 276 | 246 | 247.492 | 275 | 238 | 242 | 194 | 254 | 251 | 250 | 262 | 256 | 252 | 254.4837 | 257 | 283 | 249 | 254 | 280 | 251 | 257 | 247 | 260 | 260 | 245 | 226.3333 | 240.628 | 252.962
| Mean |  | 255.535 | 254 | 276.1667 | 252.8333 | 252.1291 | 269.1667 | 255.3333 | 242.5 | 190.1667 | 256 | 251.5 | 250.8333 | 265.1667 | 248.8333 | 254.8333 | 254.7614 | 255 | 293.1667 | 253.1667 | 249.6667 | 274.6667 | 254.3333 | 242.6667 | 243.8333 | 262.6667 | 260.5 | 245 | 236.5556 | 241.0942 | 251.7501
| Median |  | 255.91 | 254.5 | 276.5 | 254 | 252.3471 | 267.5 | 258 | 242 | 192 | 256 | 251.5 | 252.5 | 265 | 245.5 | 255 | 255.514 | 256 | 291 | 252 | 249.5 | 274 | 253.5 | 244 | 245.5 | 261.5 | 260.5 | 245.5 | 236.6667 | 240.836 | 252.6323
| Std Dev. |  | 1.2853 | 5.0596 | 4.5789 | 4.1673 | 4.8725 | 4.916 | 8.756 | 2.2583 | 4.6224 | 1.4142 | 8.735 | 7.3598 | 6.5549 | 6.4627 | 5.7067 | 2.8342 | 2.4495 | 8.9536 | 5.4191 | 3.3862 | 6.0222 | 6.5625 | 10.6333 | 4.446 | 3.7771 | 4.0373 | 3.0984 | 6.5308 | 2.2538 | 2.3565
| Rel.Std.Dev. |  | 0.503% | 1.992% | 1.658% | 1.648% | 1.933% | 1.826% | 3.429% | 0.9313% | 2.431% | 0.5524% | 3.473% | 2.934% | 2.472% | 2.597% | 2.239% | 1.113% | 0.9606% | 3.054% | 2.141% | 1.356% | 2.193% | 2.58% | 4.382% | 1.823% | 1.438% | 1.55% | 1.265% | 2.761% | 0.9348% | 0.936%
| PDM3 |  | 1.195% | 0.5869% | 9.365% | 0.1248% | -0.154% | 6.593% | 1.115% | -3.967% | -24.69% | 1.379% | -0.4032% | -0.6672% | 5.009% | -1.459% | 0.9169% | 0.8884% | 0.9829% | 16.1% | 0.2569% | -1.129% | 8.771% | 0.7189% | -3.901% | -3.439% | 4.019% | 3.161% | -2.977% | -6.321% | -4.524% | -0.3041%
| Z-Score (Absolute) |  | 0.3 | 0.06 | 3.51 | 0.12 | 0.23 | 2.42 | 0.27 | 1.72 | 9.87 | 0.38 | 0.32 | 0.43 | 1.8 | 0.74 | 0.19 | 0.18 | 0.22 | 6.16 | 0.06 | 0.61 | 3.28 | 0.12 | 1.7 | 1.52 | 1.41 | 1.08 | 1.34 | 2.65 | 1.94 | 0.29
| Analytical results for W in OREAS 45f (Certified Value 1.27 ppm) | [32 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*MS | 4A*MS |  | 2
2
| 1 | 1 | 1.1786 | 1.2 | 1.5 | 1.4 | 1.1646 | 1 | 1.4 | 1.165 | 1.8 | < 10 | 1.1 | < 0.5 | 1.2 | 1.2 | 1.2 | 1.2767 | 1.35 | 1.5 | 1.37 | 1.145 | 1.5 | 1.38 | 1.4 | 1.3 | 1.31 | 1.5 | < 10 | 1.1333 | 1.019 | 1.4086 |  | 1
| 1 | 2 | 1.1824 | 1.3 | 1.4 | 1.5 | 1.4215 | 1 | 1.44 | 1.215 | 2 | < 10 | 1.1 | < 0.5 | 1.1 | 1.7 | 1.2 | 1.2235 | 1.3 | 1.9 | 1.385 | 1.215 | 1.4 | 1.305 | 1.4 | 1.4 | 1.295 | 1.2 | < 10 | 1.2333 | 1.107 | 1.5056 |  | 11
| 1 | 3 | 1.1931 | 1.1 | 1.4 | 1.3 | 1.3765 | 1 | 1.43 | 1.14 | 1.8 | < 10 | 1.1 | < 0.5 | 1.1 | 1.3 | 1.1 | 1.3709 | 1.48 | 1.9 | 1.385 | 1.195 | 1.4 | 1.25 | 1.3 | 1.3 | 1.27 | 1.2 | < 10 | 1.1667 | 1.153 | 1.4107 |  | 16
| 1 | 4 | 1.1817 | 1.3 | 1.6 | 1.5 | 1.2844 | 1 | 1.38 | 1.18 | 1.8 | < 10 | 1.2 | 0.5 | 1.1 | 1.3 | 1.3 | 1.2562 | 1.24 | 1.5 | 1.28 | 1.25 | 1.5 | 1.25 | 1.4 | 1.4 | 1.225 | 1.3 | < 10 | 1.2 | 0.896 | 1.4175 |  | 1.2688
| 1 | 5 | 1.1843 | 1.2 | 1.5 | 1.3 | 1.4805 | 1 | 1.42 | 1.16 | 1.4 | < 10 | 1.1 | < 0.5 | 1.2 | 1.3 | 1.1 | 1.3216 | 1.41 | 2.2 | 1.16 | 1.225 | 1.5 | 1.305 | 1.3 | 1.4 | 1.3 | 1.3 | < 10 | 1.2 | 0.905 | 1.4193 |  | 133
| 1 | 6 | 1.1922 | 1.2 | 1.6 | 1.4 | 1.3892 | 1 | 1.3 | 1.12 | 2 | < 10 | 1.1 | < 0.5 | 1 | 1.2 | 1.1 | 1.309 | 1.32 | 1.3 | 1.255 | 1.28 | 1.4 | 1.265 | 1.4 | 1.4 | 1.27 | 1.2 | < 10 | 1.2 | 1.003 | 1.4467
| Mean |  | 1.1854 | 1.2167 | 1.5 | 1.4 | 1.3528 | 1 | 1.395 | 1.1633 | 1.8 |  | 1.1167 | 0.5 | 1.1167 | 1.3333 | 1.1667 | 1.293 | 1.35 | 1.7167 | 1.3058 | 1.2183 | 1.45 | 1.2925 | 1.3667 | 1.3667 | 1.2783 | 1.2833 |  | 1.1889 | 1.0138 | 1.4347
| Median |  | 1.1833 | 1.2 | 1.5 | 1.4 | 1.3829 | 1 | 1.41 | 1.1625 | 1.8 |  | 1.1 | 0.5 | 1.1 | 1.3 | 1.15 | 1.2929 | 1.335 | 1.7 | 1.325 | 1.22 | 1.45 | 1.285 | 1.4 | 1.4 | 1.2825 | 1.25 |  | 1.2 | 1.011 | 1.4184
| Std Dev. |  | 0.0059516 | 0.075277 | 0.089443 | 0.089443 | 0.11227 | 0 | 0.051284 | 0.032813 | 0.21909 |  | 0.040825 |  | 0.075277 | 0.18619 | 0.08165 | 0.052149 | 0.084853 | 0.33714 | 0.090742 | 0.04644 | 0.054772 | 0.049674 | 0.05164 | 0.05164 | 0.030768 | 0.1169 |  | 0.034427 | 0.10385 | 0.037312
| Rel.Std.Dev. |  | 0.5021% | 6.187% | 5.963% | 6.389% | 8.299% | 0% | 3.676% | 2.821% | 12.17% |  | 3.656% |  | 6.741% | 13.96% | 6.999% | 4.033% | 6.285% | 19.64% | 6.949% | 3.812% | 3.777% | 3.843% | 3.779% | 3.779% | 2.407% | 9.109% |  | 2.896% | 10.24% | 2.601%
| PDM3 |  | -6.575% | -4.109% | 18.22% | 10.34% | 6.619% | -21.19% | 9.946% | -8.313% | 41.87% |  | -11.99% | -60.59% | -11.99% | 5.086% | -8.05% | 1.905% | 6.399% | 35.3% | 2.918% | -3.978% | 14.28% | 1.868% | 7.713% | 7.713% | 0.7511% | 1.145% |  | -6.298% | -20.1% | 13.08%
| Z-Score (Absolute) |  | 0.68 | 0.5 | 1.2 | 0.6 | 0.32 | 1.8 | 0.57 | 0.82 | 3.01 | 22.23 | 1.1 | 6.05 | 1.1 | 0.2 | 0.8 | 0.04 | 0.3 | 2.51 | 0.04 | 0.49 | 0.9 | 0.04 | 0.4 | 0.4 | 0.13 | 0.1 | 22.23 | 0.66 | 1.72 | 0.81
| Analytical results for Y in OREAS 45f (Certified Value 10.9 ppm) | [32 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*OES |  | 1
1
| 1 | 1 | 10.252 | 11 | 11.8 | 10.1 | 11.3409 | 10 | 11.37 | 9.16 | 12.7 | 9 | 13.6 | 10 | 10.5 | 11.6 | 10.87 | 10.6768 | 10.1 | 9.1 | 14 | 11.25 | 12.9 | 11.5 | 11.1 | 8.7 | 11.5 | 12 | 10.5 | 10.4167 | 9.076 | 14.1608 |  | 1
| 1 | 2 | 10.279 | 10.5 | 11.9 | 10.8 | 14.3198 | 10.5 | 12.85 | 9.84 | 12.9 | 5 | 14 | 10 | 10.2 | 12.7 | 11.28 | 11.4386 | 10.2 | 9.2 | 11.75 | 11.6 | 11.6 | 11.4 | 10.8 | 9.1 | 11.55 | 11.6 | 9.9 | 10.9433 | 9.241 | 14.2857 |  | #N/A
| 1 | 3 | 10.433 | 10.1 | 11.7 | 10.4 | 12.5772 | 10 | 11.86 | 8.95 | 12.5 | 7 | 13.7 | 10 | 10.1 | 10.6 | 10.45 | 11.0031 | 10 | 9.3 | 11.55 | 10.55 | 12.1 | 11.2 | 10.5 | 8.1 | 11.75 | 11.6 | 9.6 | 10.4 | 9.29 | 14.5052 |  | 16
| 1 | 4 | 10.478 | 10.6 | 12.1 | 10.5 | 14.8127 | 10.7 | 11.56 | 10 | 12.8 | 6 | 13.8 | 10 | 10.3 | 12.1 | 10.82 | 11.1888 | 10.2 | 9.2 | 11.35 | 10.45 | 12.1 | 11.1 | 11.3 | 7.9 | 11.9 | 11 | 10.1 | 10.5967 | 9.06 | 13.8578 |  | 10.8951
| 1 | 5 | 10.257 | 10.1 | 12 | 10.2 | 13.3032 | 10.1 | 11.54 | 8.68 | 12.2 | 7 | 12.2 | 10 | 10.8 | 12.1 | 10.57 | 11.1306 | 10.5 | 9.1 | 10.05 | 10.95 | 12.5 | 9.67 | 11.1 | 8.9 | 12.3 | 11.2 | 9.5 | 10.76 | 9.299 | 14.2039 |  | 134
| 1 | 6 | 10.34 | 10.2 | 11.1 | 10.8 | 12.786 | 10.2 | 11.72 | 7.67 | 12.7 | 6 | 11.4 | 10 | 10.5 | 11.1 | 10.51 | 11.2445 | 10.3 | 9 | 10.8 | 10.35 | 11.6 | 10.1 | 11 | 8.8 | 11.5 | 11.2 | 10 | 10.52 | 9.045 | 13.3759
| Mean |  | 10.3398 | 10.4167 | 11.7667 | 10.4667 | 13.19 | 10.25 | 11.8167 | 9.05 | 12.6333 | 6.6667 | 13.1167 | 10 | 10.4 | 11.7 | 10.75 | 11.1137 | 10.2167 | 9.15 | 11.5833 | 10.8583 | 12.1333 | 10.8283 | 10.9667 | 8.5833 | 11.75 | 11.4333 | 9.9333 | 10.6061 | 9.1685 | 14.0649
| Median |  | 10.3095 | 10.35 | 11.85 | 10.45 | 13.0446 | 10.15 | 11.64 | 9.055 | 12.7 | 6.5 | 13.65 | 10 | 10.4 | 11.85 | 10.695 | 11.1597 | 10.2 | 9.15 | 11.45 | 10.75 | 12.1 | 11.15 | 11.05 | 8.75 | 11.65 | 11.4 | 9.95 | 10.5583 | 9.1585 | 14.1823
| Std Dev. |  | 0.095965 | 0.35449 | 0.3559 | 0.29439 | 1.2554 | 0.2881 | 0.53309 | 0.84735 | 0.25033 | 1.3663 | 1.0591 | 0 | 0.25298 | 0.76158 | 0.3099 | 0.25757 | 0.17224 | 0.10488 | 1.3333 | 0.4964 | 0.50859 | 0.75658 | 0.28048 | 0.47504 | 0.31305 | 0.36697 | 0.36148 | 0.21132 | 0.12052 | 0.39716
| Rel.Std.Dev. |  | 0.9281% | 3.403% | 3.025% | 2.813% | 9.518% | 2.811% | 4.511% | 9.363% | 1.982% | 20.49% | 8.074% | 0% | 2.433% | 6.509% | 2.883% | 2.318% | 1.686% | 1.146% | 11.51% | 4.572% | 4.192% | 6.987% | 2.558% | 5.534% | 2.664% | 3.21% | 3.639% | 1.992% | 1.315% | 2.824%
| PDM3 |  | -5.096% | -4.391% | 8% | -3.932% | 21.06% | -5.921% | 8.459% | -16.93% | 15.95% | -38.81% | 20.39% | -8.215% | -4.544% | 7.388% | -1.332% | 2.007% | -6.227% | -16.02% | 6.317% | -0.3372% | 11.37% | -0.6126% | 0.6571% | -21.22% | 7.847% | 4.94% | -8.827% | -2.652% | -15.85% | 29.09%
| Z-Score (Absolute) |  | 0.38 | 0.32 | 0.73 | 0.28 | 1.83 | 0.45 | 0.76 | 1.38 | 1.4 | 3.22 | 1.77 | NA | 0.33 | 0.67 | 0.06 | 0.22 | 0.47 | 1.3 | 0.58 | 0.02 | 1.01 | 0 | 0.11 | 1.74 | 0.71 | 0.47 | 0.69 | 0.17 | 1.28 | 2.5
Results from laboratory 19 were removed due to their 10 ppm reading resolution.
| Analytical results for Yb in OREAS 45f (Certified Value 1.25 ppm) | [32 empty] | Certified
Yb | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab |  | 1
| No | No | 02: 02; 04: 04; 05: 05; 07: 12; 10: 13; 12: 15; 13: 18; 14: 23; 15: 24; 17: 25; 18: 27; 19: 33; 20: 34; 21: 38; 23: 39; 24: 40 |  | ppm
|  |  | 02: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 07: 4A*MS; 10: 4A*OES/MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*MS; 15: 4A*OES/MS; 17: 4A*MS; 18: 4A*OES/MS; 19: 4A*MS; 20: 4A*MS; 21: 4A*MS; 23: 5A*MS; 24: 4A*MS |  | 2
2
| 1 | 1 | 02: 1.2637; 04: 1.1; 05: 1.4; 07: 1.4; 10: 1.3; 12: 1.1; 13: 1.4; 14: 1.23; 15: 1.291; 17: 1.1; 18: 1.3; 19: 1.4; 20: 1; 21: 1.27; 23: 1.088; 24: 1.6682 |  | 1
| 1 | 2 | 02: 1.2526; 04: 1.2; 05: 1.3; 07: 1.35; 10: 1.56; 12: 1.2; 13: 1.3; 14: 1.19; 15: 1.3362; 17: 1.14; 18: 1.3; 19: 1.3; 20: 1.1; 21: 1.2967; 23: 1.017; 24: 1.7482 |  | #N/A
| 1 | 3 | 02: 1.2578; 04: 1.1; 05: 1.4; 07: 1.2; 10: 1.35; 12: 1.4; 13: 1.3; 14: 1.25; 15: 1.2941; 17: 1.07; 18: 1.3; 19: 1.4; 20: 1; 21: 1.2333; 23: 1.028; 24: 1.6537 |  | 16
| 1 | 4 | 02: 1.2361; 04: 1.2; 05: 1.4; 07: 1.35; 10: 1.29; 12: 1.3; 13: 1.6; 14: 1.17; 15: 1.2786; 17: 1.02; 18: 1.3; 19: 1.5; 20: 1; 21: 1.2567; 23: 1.021; 24: 1.6765 |  | 1.2476
| 1 | 5 | 02: 1.2619; 04: 1.2; 05: 1.4; 07: 1.2; 10: 1.44; 12: 1.1; 13: 1.4; 14: 1.27; 15: 1.3356; 17: 1.1; 18: 1.2; 19: 1.5; 20: 1; 21: 1.2733; 23: 1.023; 24: 1.6659 |  | 135
| 1 | 6 | 02: 1.2399; 04: 1.2; 05: 1.3; 07: 1.2; 10: 1.37; 12: 1.4; 13: 1.3; 14: 1.28; 15: 1.2918; 17: 1.09; 18: 1.2; 19: 1.5; 20: 1; 21: 1.2733; 23: 1.012; 24: 1.6085
| Mean |  | 02: 1.252; 04: 1.1667; 05: 1.3667; 07: 1.2833; 10: 1.385; 12: 1.25; 13: 1.3833; 14: 1.2317; 15: 1.3045; 17: 1.0867; 18: 1.2667; 19: 1.4333; 20: 1.0167; 21: 1.2672; 23: 1.0315; 24: 1.6702
| Median |  | 02: 1.2552; 04: 1.2; 05: 1.4; 07: 1.275; 10: 1.36; 12: 1.25; 13: 1.35; 14: 1.24; 15: 1.2929; 17: 1.095; 18: 1.3; 19: 1.45; 20: 1; 21: 1.2717; 23: 1.022; 24: 1.667
| Std Dev. |  | 02: 0.01156; 04: 0.05164; 05: 0.05164; 07: 0.093095; 10: 0.10134; 12: 0.13784; 13: 0.1169; 14: 0.044008; 15: 0.024886; 17: 0.039833; 18: 0.05164; 19: 0.08165; 20: 0.040825; 21: 0.02102; 23: 0.028205; 24: 0.045222
| Rel.Std.Dev. |  | 02: 0.9233%; 04: 4.426%; 05: 3.779%; 07: 7.254%; 10: 7.317%; 12: 11.03%; 13: 8.451%; 14: 3.573%; 15: 1.908%; 17: 3.666%; 18: 4.077%; 19: 5.696%; 20: 4.016%; 21: 1.659%; 23: 2.734%; 24: 2.708%
| PDM3 |  | 02: 0.3528%; 04: -6.487%; 05: 9.544%; 07: 2.864%; 10: 11.01%; 12: 0.1925%; 13: 10.88%; 14: -1.277%; 15: 4.565%; 17: -12.9%; 18: 1.528%; 19: 14.89%; 20: -18.51%; 21: 1.573%; 23: -17.32%; 24: 33.87%
| Z-Score (Absolute) |  | 02: 0.1; 04: 0.68; 05: 0.67; 07: 0.11; 10: 0.8; 12: 0.11; 13: 0.78; 14: 0.24; 15: 0.25; 17: 1.22; 18: 0; 19: 1.12; 20: 1.69; 21: 0; 23: 1.59; 24: 2.72
| Analytical results for Zn in OREAS 45f (Certified Value 35.3 ppm) | [32 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 5A*OES | 4A*MS |  | 1
2
| 1 | 1 | 36.7679 | 37.5 | 36 | 49 | 48.2085 | 34 | 37 | 35.8 | 42 | 47 | 58 | 35 | 33 | 30 | 36 | 34.3537 | 34.7 | 35.6171 | 34 | 34.3 | 35 | 36.1 | 37 | 46 | 36.6 | 32.1 | 35 | 33.6667 | 29.362 | 41.6298 |  | 1
| 1 | 2 | 37.3657 | 38.7 | 35 | 48 | 47.1256 | 34 | 37 | 34.6 | 48 | 52 | 54 | 35 | 32 | 31 | 36 | 36.0528 | 33.5 | 33.1335 | 35.1 | 35 | 37 | 36.4 | 37 | 46 | 36.3 | 33.2 | 36 | 35 | 29.61 | 38.1259 |  | 1
| 1 | 3 | 37.4116 | 38.5 | 38 | 47 | 46.8191 | 34 | 39 | 33.6 | 42 | 46 | 56 | 35 | 34 | 30 | 36 | 34.4872 | 34.4 | 35.2703 | 35.3 | 35 | 34 | 35.3 | 36 | 44 | 35.9 | 31.8 | 34 | 33 | 29.922 | 41.5073 |  | 16
| 1 | 4 | 36.2294 | 39 | 36 | 47 | 48.5555 | 36 | 38 | 34.6 | 46 | 47 | 55 | 35 | 32 | 30 | 37 | 36.1536 | 34.8 | 33.1286 | 35.7 | 34.7 | 39 | 35.4 | 35 | 44 | 36.3 | 32.3 | 35 | 33.6667 | 29.213 | 41.1701 |  | 35.2584
| 1 | 5 | 36.365 | 39.6 | 37 | 47 | 46.0082 | 32 | 34 | 35 | 44 | 48 | 54 | 35 | 34 | 30 | 36 | 35.4498 | 32.9 | 33.8146 | 33.6 | 33.5 | 38 | 34.1 | 36 | 45 | 35.1 | 35 | 34 | 33.6667 | 29.103 | 39.9206 |  | 136
| 1 | 6 | 37.5758 | 36.3 | 39 | 48 | 48.4162 | 36 | 35 | 33 | 47 | 47 | 54 | 35 | 32 | 30 | 37 | 36.4954 | 34.4 | 33.7656 | 34.6 | 35.9 | 37 | 34.3 | 40 | 45 | 35.7 | 32.2 | 34 | 34.6667 | 28.464 | 39.1815
| Mean |  | 36.9526 | 38.2667 | 36.8333 | 47.6667 | 47.5222 | 34.3333 | 36.6667 | 34.4333 | 44.8333 | 47.8333 | 55.1667 | 35 | 32.8333 | 30.1667 | 36.3333 | 35.4988 | 34.1167 | 34.1216 | 34.7167 | 34.7333 | 36.6667 | 35.2667 | 36.8333 | 45 | 35.9833 | 32.7667 | 34.6667 | 33.9444 | 29.279 | 40.2559
| Median |  | 37.0668 | 38.6 | 36.5 | 47.5 | 47.6671 | 34 | 37 | 34.6 | 45 | 47 | 54.5 | 35 | 32.5 | 30 | 36 | 35.7513 | 34.4 | 33.7901 | 34.85 | 34.85 | 37 | 35.35 | 36.5 | 45 | 36.1 | 32.25 | 34.5 | 33.6667 | 29.2875 | 40.5453
| Std Dev. |  | 0.57827 | 1.1843 | 1.472 | 0.8165 | 1.0278 | 1.5055 | 1.8619 | 0.99933 | 2.5626 | 2.137 | 1.6021 | 0 | 0.98319 | 0.40825 | 0.5164 | 0.90176 | 0.75211 | 1.0714 | 0.80353 | 0.80166 | 1.8619 | 0.92664 | 1.7224 | 0.89443 | 0.53821 | 1.1911 | 0.8165 | 0.74287 | 0.49587 | 1.4208
| Rel.Std.Dev. |  | 1.565% | 3.095% | 3.996% | 1.713% | 2.163% | 4.385% | 5.078% | 2.902% | 5.716% | 4.468% | 2.904% | 0% | 2.994% | 1.353% | 1.421% | 2.54% | 2.205% | 3.14% | 2.315% | 2.308% | 5.078% | 2.628% | 4.676% | 1.988% | 1.496% | 3.635% | 2.355% | 2.188% | 1.694% | 3.529%
| PDM3 |  | 4.805% | 8.532% | 4.467% | 35.19% | 34.78% | -2.624% | 3.994% | -2.34% | 27.16% | 35.67% | 56.46% | -0.7329% | -6.878% | -14.44% | 3.049% | 0.6817% | -3.238% | -3.224% | -1.536% | -1.489% | 3.994% | 0.02344% | 4.467% | 27.63% | 2.056% | -7.067% | -1.678% | -3.727% | -16.96% | 14.17%
| Z-Score (Absolute) |  | 0.54 | 1.13 | 0.49 | 5.31 | 5.25 | 0.63 | 0.41 | 0.58 | 4.05 | 5.39 | 8.65 | 0.33 | 1.3 | 2.48 | 0.26 | 0.11 | 0.72 | 0.72 | 0.46 | 0.45 | 0.41 | 0.21 | 0.49 | 4.12 | 0.11 | 1.33 | 0.48 | 0.8 | 2.88 | 2.01
| Analytical results for Zr in OREAS 45f (Certified Value 172 ppm) | [32 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 5A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 164.8224 | 174.7 | 179 | 176 | 206.9452 | 184 | 166.8 | 173 | 175 | 161 | 201 | 130 | 165.4 | 159.3 | 171.6 | 169.3934 | 161 | 180.5 | 169.5 | 191.5 | 181 | 193 | 154.4 | 181 | 239 | 163 | 166.3333 | 155.491 | 227.494 |  |  | 1
| 1 | 2 | 167.6548 | 178.6 | 172 | 189 | 191.7429 | 194 | 178.4 | 169.5 | 181 | 147 | 193 | 115 | 165.2 | 165.8 | 172.6 | 170.0131 | 167 | 178.5 | 171.5 | 184 | 179.5 | 157 | 154.7 | 191 | 239 | 157 | 173.6333 | 159.629 | 226.651 |  |  | #N/A
| 1 | 3 | 164.1143 | 169.6 | 178 | 182 | 189.859 | 170 | 166.1 | 165.5 | 166 | 152 | 196 | 130 | 166 | 163.2 | 168.5 | 169.3416 | 159 | 178 | 168.5 | 179.5 | 182.5 | 156 | 149.7 | 185 | 229 | 155 | 165.5333 | 160.9 | 234.8724 |  |  | 16
| 1 | 4 | 169.9828 | 178.5 | 178 | 179 | 194.8514 | 194 | 163.3 | 167.5 | 177 | 149 | 198 | 130 | 173.4 | 169.6 | 175 | 173.2052 | 162 | 168 | 168 | 185.5 | 176.5 | 161 | 155 | 185 | 226 | 161 | 167.2667 | 158.826 | 228.7284 |  |  | 171.942
| 1 | 5 | 166.0349 | 168.2 | 186 | 176 | 179.9428 | 175 | 163.1 | 167.5 | 165 | 155 | 190 | 120 | 168.6 | 159.7 | 172.5 | 176.5447 | 166 | 178 | 172 | 182.5 | 167 | 169 | 154.2 | 183.5 | 226 | 154 | 169.3333 | 158.051 | 227.4446 |  |  | 137
| 1 | 6 | 165.967 | 169.3 | 181 | 182 | 179.5026 | 188 | 176.8 | 165.5 | 176 | 149 | 191 | 120 | 170.4 | 166.2 | 168.3 | 175.3951 | 166 | 180 | 174.5 | 185.5 | 168 | 177 | 156.1 | 181.5 | 218 | 162 | 163.4333 | 157.487 | 234.6219
| Mean |  | 166.4294 | 173.15 | 179 | 180.6667 | 190.474 | 184.1667 | 169.0833 | 168.0833 | 173.3333 | 152.1667 | 194.8333 | 124.1667 | 168.1667 | 163.9667 | 171.4167 | 172.3155 | 163.5 | 177.1667 | 170.6667 | 184.75 | 175.75 | 168.8333 | 154.0167 | 184.5 | 229.5 | 158.6667 | 167.5889 | 158.3973 | 229.9687
| Median |  | 166.0009 | 172.15 | 178.5 | 180.5 | 190.8009 | 186 | 166.45 | 167.5 | 175.5 | 150.5 | 194.5 | 125 | 167.3 | 164.5 | 172.05 | 171.6091 | 164 | 178.25 | 170.5 | 184.75 | 178 | 165 | 154.55 | 184.25 | 227.5 | 159 | 166.8 | 158.4385 | 228.1112
| Std Dev. |  | 2.1185 | 4.7467 | 4.5607 | 4.8854 | 10.2345 | 9.9281 | 6.778 | 2.8358 | 6.4083 | 5.1543 | 4.2622 | 6.6458 | 3.2776 | 4.0163 | 2.5949 | 3.1889 | 3.2711 | 4.6116 | 2.4631 | 3.9969 | 6.6989 | 14.2607 | 2.2176 | 3.6056 | 8.2158 | 3.8297 | 3.5427 | 1.8633 | 3.7614
| Rel.Std.Dev. |  | 1.273% | 2.741% | 2.548% | 2.704% | 5.373% | 5.391% | 4.009% | 1.687% | 3.697% | 3.387% | 2.188% | 5.352% | 1.949% | 2.449% | 1.514% | 1.851% | 2.001% | 2.603% | 1.443% | 2.163% | 3.812% | 8.447% | 1.44% | 1.954% | 3.58% | 2.414% | 2.114% | 1.176% | 1.636%
| PDM3 |  | -3.206% | 0.7026% | 4.105% | 5.074% | 10.78% | 7.11% | -1.663% | -2.244% | 0.8092% | -11.5% | 13.31% | -27.79% | -2.196% | -4.638% | -0.3055% | 0.2173% | -4.91% | 3.039% | -0.7417% | 7.449% | 2.215% | -1.808% | -10.43% | 7.304% | 33.48% | -7.721% | -2.532% | -7.877% | 33.75%
| Z-Score (Absolute) |  | 0.44 | 0.15 | 0.67 | 0.82 | 1.69 | 1.13 | 0.21 | 0.3 | 0.17 | 1.71 | 2.08 | 4.2 | 0.29 | 0.66 | 0 | 0.08 | 0.7 | 0.51 | 0.07 | 1.19 | 0.39 | 0.23 | 1.55 | 1.16 | 5.16 | 1.13 | 0.34 | 1.16 | 5.21
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C37: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C258: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C276: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C294: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C312: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C566: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C584: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C602: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1024: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1062: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1080: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1098: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1117: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1135: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1171: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1189: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1208: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1226: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1244: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 0.15-50g
| Analytical results for Ag in OREAS 45f (Indicative Value 0.057 ppm) | [31 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 3
| 1 | 1 | < 0.5 | 0.07 | 0.05 | < 0.05 | 0.52424 | 0.06 | 0.053 | 0.24 | < 0.5 | < 0.1 | 0.05 | 0.066 | < 0.1 | 0.05 | 0.054249 | 0.069 | 0.09 | 0.045 | 0.045 | 0.06 | 0.046 | < 0.1 | 0.47 | 0.041 | 0.1 | < 1 | 0.07 | 0.22 | 0.04972 |  | 1
| 1 | 2 | < 0.5 | 0.08 | 0.05 | 0.05 | 0.66207 | 0.05 | 0.052 | 0.22 | < 0.5 | < 0.1 | 0.05 | 0.062 | < 0.1 | < 0.05 | 0.054934 | 0.0666 | 0.1 | 0.052 | 0.046 | 0.05 | 0.047 | < 0.1 | 0.4 | 0.047 | 0.11 | < 1 | 0.06 | 0.19 | 0.04837 |  | 7
| 1 | 3 | < 0.5 | 0.1 | 0.06 | < 0.05 | 0.62145 | 0.06 | 0.053 | 0.22 | < 0.5 | 0.1 | 0.05 | 0.059 | < 0.1 | < 0.05 | 0.050774 | 0.0641 | 0.09 | 0.048 | 0.049 | 0.05 | 0.048 | < 0.1 | 0.39 | 0.046 | 0.08 | < 1 | 0.05 | 0.221 | 0.04846 |  | 16
| 1 | 4 | < 0.5 | 0.09 | 0.05 | < 0.05 | 0.51858 | 0.06 | 0.052 | 0.22 | < 0.5 | < 0.1 | 0.05 | 0.056 | < 0.1 | < 0.05 | 0.070457 | 0.0646 | 0.09 | 0.048 | 0.05 | 0.05 | 0.048 | < 0.1 | 0.43 | 0.047 | 0.1 | < 1 | 0.06 | 0.228 | 0.04813 |  | 0.05665
| 1 | 5 | < 0.5 | 0.09 | 0.05 | 0.05 | 0.64528 | 0.06 | 0.052 | 0.26 | < 0.5 | 0.1 | 0.05 | 0.053 | < 0.1 | < 0.05 | 0.066011 | 0.0647 | 0.09 | 0.047 | 0.046 | 0.06 | 0.046 | < 0.1 | 0.42 | 0.052 | 0.1 | < 1 | 0.065 | 0.219 | 0.05077 |  | 27
| 1 | 6 | < 0.5 | 0.09 | 0.05 | 0.05 | 0.63453 | 0.04 | 0.047 | 0.23 | < 0.5 | 0.2 | 0.05 | 0.06 | < 0.1 | 0.05 | 0.057444 | 0.0708 | 0.09 | 0.045 | 0.045 | 0.05 | 0.049 | < 0.1 | 0.41 | 0.047 | 0.1 | < 1 | 0.06 | 0.208 | 0.04972
| Mean |  |  | 0.086667 | 0.051667 | 0.05 | 0.60102 | 0.055 | 0.0515 | 0.23167 |  | 0.13333 | 0.05 | 0.059333 |  | 0.05 | 0.058978 | 0.066633 | 0.091667 | 0.0475 | 0.046833 | 0.053333 | 0.047333 |  | 0.42 | 0.046667 | 0.098333 |  | 0.060833 | 0.21433 | 0.049195
| Median |  |  | 0.09 | 0.05 | 0.05 | 0.62799 | 0.06 | 0.052 | 0.225 |  | 0.1 | 0.05 | 0.0595 |  | 0.05 | 0.056189 | 0.06565 | 0.09 | 0.0475 | 0.046 | 0.05 | 0.0475 |  | 0.415 | 0.047 | 0.1 |  | 0.06 | 0.2195 | 0.04909
| Std Dev. |  |  | 0.010328 | 0.0040825 | 8.4984e-18 | 0.063117 | 0.0083666 | 0.0022583 | 0.016021 |  | 0.057735 | 7.6012e-18 | 0.0045461 |  | 0 | 0.0076106 | 0.0027296 | 0.0040825 | 0.0025884 | 0.002137 | 0.005164 | 0.0012111 |  | 0.028284 | 0.0035024 | 0.0098319 |  | 0.0066458 | 0.013545 | 0.001038
| Rel.Std.Dev. |  |  | 11.92% | 7.902% | 1.7e-14% | 10.5% | 15.21% | 4.385% | 6.915% |  | 43.3% | 1.52e-14% | 7.662% |  | 0% | 12.9% | 4.096% | 4.454% | 5.449% | 4.563% | 9.682% | 2.559% |  | 6.734% | 7.505% | 9.999% |  | 10.92% | 6.32% | 2.11%
| PDM3 |  |  | 52.99% | -8.797% | -11.74% | 960.9% | -2.913% | -9.091% | 308.9% |  | 135.4% | -11.74% | 4.736% |  | -11.74% | 4.109% | 17.62% | 61.81% | -16.15% | -17.33% | -5.855% | -16.45% |  | 641.4% | -17.62% | 73.58% |  | 7.384% | 278.3% | -13.16%
| Z-Score (Absolute) |  | 10.61 | 1.54 | 0.41 | 1.19 | 30.1 | 0.22 | 0.42 | 9.59 | 10.61 | 1.81 | 0.5 | 0.02 | 0.5 | 1.42 | 0 | 0.43 | 1.81 | 0.64 | 0.67 | 0.31 | 0.65 | 0.5 | 20.04 | 0.68 | 2.19 | 24.49 | 0.1 | 8.63 | 0.54
| Analytical results for Al in OREAS 45f (Certified Value 4.81 wt.%) | [31 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 6.3774 | 4.83 | 4.79 | 7.74 | 7.2 | 4.81 | 4.72 | 6.18 | 3.24 | 4.79 | 4.65 | 3.96 | 4.853 | 6.9869 | 4.75 | 4.02 | 5.66 | 5.25 | 4.93 | 4.78 | > 5 | 3.9 | 4.8 | 6.19 | 3.66 | 7.7953 | 1.2649 | 5.0411 |  |  | 1
| 1 | 2 | 6.3675 | 4.83 | 5.01 | 7.87 | 7.1 | 4.86 | 4.4859 | 6.11 | 3.14 | 4.55 | 4.71 | 4.12 | 5.203 | 6.9639 | 4.54 | 4.15 | 5.79 | 5.12 | 5.03 | 4.87 |  | 3.48 | 4.63 | 5.14 | 3.66 | 7.4022 | 1.1999 | 4.7799 |  |  | #N/A
| 1 | 3 | 6.3842 | 5.01 | 5.35 | 7.7 | 6.86 | 5.02 | 4.7797 | 6.13 | 3.09 | 4.71 | 4.72 | 4.11 | 5 | 7.2847 | 5.37 | 3.94 | 5.94 | 5.3 | 5.08 | 4.87 | 4.94 | 3.15 | 5.02 | 4.9 | 3.69 | 7.3311 | 1.3364 | 5.2013 |  |  | 16
| 1 | 4 | 6.3822 | 4.84 | 5.33 | 7.71 | 7.34 | 4.95 | 4.8215 | 6.15 | 3.24 | 4.42 | 4.67 | 4.15 | 4.631 | 7.5033 | 5.04 | 4.14 | 5.86 | 5.23 | 5.2 | 4.57 | > 5 | 3.63 | 5.01 | 4.92 | 3.63 | 7.6714 | 1.2547 | 5.2876 |  |  | 4.8054
| 1 | 5 | 6.3651 | 4.94 | 5.27 | 7.79 | 6.72 | 4.95 | 4.4957 | 6.13 | 3 | 4.07 | 4.62 | 4.07 | 4.733 | 7.3819 | 5.31 | 3.81 | 5.92 | 5.15 | 4.79 | 4.69 |  | 3.82 | 4.84 | 4.62 | 3.67 | 7.3691 | 1.3225 | 5.405 |  |  | 139
| 1 | 6 | 6.3544 | 4.93 | 5.23 | 7.65 | 6.52 | 4.78 | 4.6746 | 6.13 | 3.04 | 4.48 | 4.77 | 4.1 | 4.859 | 7.2162 | 4.76 | 4.22 | 5.72 | 5.13 | 5.39 | 4.55 |  | 3.87 | 4.83 | 4.97 | 3.57 | 7.2596 | 1.3297 | 5.2951
| Mean |  | 6.3718 | 4.8967 | 5.1633 | 7.7433 | 6.9567 | 4.895 | 4.6629 | 6.1383 | 3.125 | 4.5033 | 4.69 | 4.085 | 4.8798 | 7.2228 | 4.9617 | 4.0467 | 5.815 | 5.1967 | 5.07 | 4.7217 | 4.94 | 3.6417 | 4.855 | 5.1233 | 3.6467 | 7.4714 | 1.2847 | 5.1683
| Median |  | 6.3724 | 4.885 | 5.25 | 7.725 | 6.98 | 4.905 | 4.6973 | 6.13 | 3.115 | 4.515 | 4.69 | 4.105 | 4.856 | 7.2504 | 4.9 | 4.08 | 5.825 | 5.19 | 5.055 | 4.735 | 4.94 | 3.725 | 4.835 | 4.945 | 3.66 | 7.3857 | 1.2937 | 5.2444
| Std Dev. |  | 0.011488 | 0.074744 | 0.2197 | 0.077374 | 0.31072 | 0.093113 | 0.14246 | 0.024014 | 0.10075 | 0.2539 | 0.054037 | 0.066558 | 0.20175 | 0.21473 | 0.33391 | 0.15332 | 0.11167 | 0.073666 | 0.20928 | 0.14204 |  | 0.28937 | 0.14543 | 0.54884 | 0.042269 | 0.21198 | 0.054061 | 0.22557
| Rel.Std.Dev. |  | 0.1803% | 1.526% | 4.255% | 0.9992% | 4.467% | 1.902% | 3.055% | 0.3912% | 3.224% | 5.638% | 1.152% | 1.629% | 4.134% | 2.973% | 6.73% | 3.789% | 1.92% | 1.418% | 4.128% | 3.008% |  | 7.946% | 2.995% | 10.71% | 1.159% | 2.837% | 4.208% | 4.364%
| PDM3 |  | 32.6% | 1.9% | 7.449% | 61.14% | 44.77% | 1.865% | -2.965% | 27.74% | -34.97% | -6.285% | -2.401% | -14.99% | 1.55% | 50.31% | 3.253% | -15.79% | 21.01% | 8.143% | 5.507% | -1.742% | 2.802% | -24.22% | 1.033% | 6.617% | -24.11% | 55.48% | -73.27% | 7.553%
| Z-Score (Absolute) |  | 2.83 | 0.04 | 0.48 | 5.49 | 3.96 | 0.05 | 0.5 | 2.37 | 3.49 | 0.81 | 0.44 | 1.62 | 0.07 | 4.48 | 0.08 | 1.7 | 1.74 | 0.54 | 0.3 | 0.38 | 0.04 | 2.48 | 0.12 | 0.4 | 2.47 | 4.97 | 7.07 | 0.49
| Analytical results for As in OREAS 45f (Indicative Value 2.7 ppm) | [31 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | ppm
|  |  | AR*OES | Various [a] | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | Various [b] | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | 2
|  |  | 0.5g | see footer | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | see footer | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 6.1561 | < 3 | 3 | 3 | 10.3087 | 1.7 | 2.59 | 4.3 | < 2 | 2 | 2.8 | 2.2 | 1.1 | 2.5 | 3.7661 | 2.61 | < 5 | 2 | 2.54 | 2.6 | 2.1 | 2.69 | 4 | 4 | 2.61 | 6 | 8 | 4 | 2.83 | 2.9389 | 1
| 1 | 2 | 6.2432 | 4 | 3 | 3 | 10.7653 | 1.7 | 2.46 | 4.3 | 3 | 2 | 2.6 | 2.5 | 1.7 | 2.3 | 3.6375 | 2.61 | < 5 | 2 | 2.63 | 2.31 | 2 | 2.62 | 2 | 3 | 2.58 | 4.9 | 7 | 4 | 3.792 | 3.2809 | 1
| 1 | 3 | 6.5707 | 4 | 3 | 3 | 10.1206 | 1.6 | 2.79 | 4.2 | 4 | 2 | 2.3 | 2.1 | 1.2 | 2.4 | 3.7354 | 2.89 | < 5 | 2 | 2.57 | 2.54 | 1.8 | 2.99 | 2 | 3 | 2.96 | 6.1 | 7 | 3.3333 | 2.673 | 3.1746 | 16
| 1 | 4 | 6.3154 | 3 | 3 | 3 | 11.2173 | 1.8 | 2.46 | 4 | 3 | 2 | 2.9 | 2.1 | 1.5 | 2.5 | 3.8709 | 2.31 | 6.2473 | 2 | 2.57 | 2.39 | 1.7 | 2.62 | 3 | 3 | 2.73 | 5.3 | 8 | 4.3333 | 2.83 | 3.3211 | 2.7
| 1 | 5 | 6.2365 | 4 | 3 | 3 | 10.4277 | 1.7 | 2.3 | 4.2 | < 2 | 2 | 2.8 | 2.1 | 1.1 | 2.4 | 3.9216 | 2.37 | 5.8486 | 2 | 2.41 | 2.39 | 2.1 | 3.03 | 3 | 3 | 2.73 | 7.4 | 7 | 3 | 3.839 | 3.0875 | 28
| 1 | 6 | 6.5638 | 3 | 3 | 3 | 10.1331 | 1.6 | 2.35 | 4.2 | 3 | 2 | 3 | 2 | 1.6 | 2.3 | 3.7535 | 2.46 | < 5 | 2 | 2.59 | 2.51 | 2.1 | 2.56 | 4 | 3 | 2.87 | 5.1 | 8 | 4 | 3.12 | 3.0484
| Mean |  | 6.3476 | 3.6 | 3 | 3 | 10.4954 | 1.6833 | 2.4917 | 4.2 | 3.25 | 2 | 2.7333 | 2.1667 | 1.3667 | 2.4 | 3.7808 | 2.5417 | 6.0479 | 2 | 2.5517 | 2.4567 | 1.9667 | 2.7517 | 3 | 3.1667 | 2.7467 | 5.8 | 7.5 | 3.7778 | 3.1807 | 3.1419
| Median |  | 6.2793 | 4 | 3 | 3 | 10.3682 | 1.7 | 2.46 | 4.2 | 3 | 2 | 2.8 | 2.1 | 1.35 | 2.4 | 3.7598 | 2.535 | 6.0479 | 2 | 2.57 | 2.45 | 2.05 | 2.655 | 3 | 3 | 2.73 | 5.65 | 7.5 | 4 | 2.975 | 3.131
| Std Dev. |  | 0.17746 | 0.54772 | 0 | 0 | 0.42549 | 0.075277 | 0.17747 | 0.10954 | 0.5 | 0 | 0.25033 | 0.17512 | 0.26583 | 0.089443 | 0.10146 | 0.20999 | 0.28192 | 0 | 0.075476 | 0.11021 | 0.17512 | 0.20469 | 0.89443 | 0.40825 | 0.14679 | 0.92087 | 0.54772 | 0.50185 | 0.51273 | 0.1452
| Rel.Std.Dev. |  | 2.796% | 15.21% | 0% | 0% | 4.054% | 4.472% | 7.123% | 2.608% | 15.38% | 0% | 9.159% | 8.082% | 19.45% | 3.727% | 2.684% | 8.262% | 4.661% | 0% | 2.958% | 4.486% | 8.904% | 7.439% | 29.81% | 12.89% | 5.344% | 15.88% | 7.303% | 13.28% | 16.12% | 4.622%
| PDM3 |  | 135.1% | 33.33% | 11.11% | 11.11% | 288.7% | -37.65% | -7.715% | 55.56% | 20.37% | -25.93% | 1.236% | -19.75% | -49.38% | -11.11% | 40.03% | -5.863% | 124% | -25.93% | -5.493% | -9.011% | -27.16% | 1.915% | 11.11% | 17.29% | 1.73% | 114.8% | 177.8% | 39.92% | 17.8% | 16.37%
| Z-Score (Absolute) |  | 4.19 | NA | NA | NA | 9.01 | 1.24 | 0.3 | 1.69 | NA | NA | 0.02 | 0.67 | 1.6 | 0.4 | 1.2 | 0.24 | 1.09 | NA | 0.23 | 0.34 | 0.91 | 0.01 | NA | NA | 0 | 3.55 | NA | 1.2 | 0.5 | 0.46
Results from laboratories 4, 5, 17, 18, 27, 33, 34 & 37 were removed due to their 1 ppm reading resolution.
[a]: Lab 04: Rnd 01 - AR*MS (0.25g), Rnd 01 - AR*OES (0.25g)
[b]: Lab 33: Rnd 01 - AR*MS (0.4g), Rnd 01 - AR*OES (0.4g)
| Analytical results for B in OREAS 45f (Indicative Value < 10 ppm) | [31 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 04 | 10 | 14 | 17 | 18 | 20 | 21 | 24 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | [12 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | [12 empty] | 1
|  |  | 0.25g | 01g | 0.5g | 0.5g | 0.4g | 15g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | [12 empty] | 1
| 1 | 1 | < 10 | 9.102 | < 10 | < 5 | 12 | 2 | < 20 | < 10 | < 10 | 10 | < 10 | < 10 | < 10 | < 10 | < 10 | 22 | < 10 | < 5 | [12 empty] | 1
| 1 | 2 | < 10 | 8.5803 | < 10 | < 5 | 13 | 2 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 5 | [12 empty] | 9
| 1 | 3 | < 10 | 8.7755 | < 10 | < 5 | 12 | 2 | < 20 | < 10 | 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 5 | [12 empty] | 16
| 1 | 4 | < 10 | 8.4791 | < 10 | < 5 | 12 | 2 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 5 | [12 empty] | < 10
| 1 | 5 | < 10 | 8.9311 | < 10 | < 5 | 11 | 2 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 5 | [12 empty] | 29
| 1 | 6 | < 10 | 8.0913 | < 10 | < 5 | 12 | 2 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 5
| Mean |  | 04: 8.6599; 10: 12; 13: 2; 17: 10; 18: 10; 25: 22
| Median |  | 04: 8.6779; 10: 12; 13: 2; 17: 10; 18: 10; 25: 22
| Std Dev. |  | 04: 0.35907; 10: 0.63246; 13: 0
| Rel.Std.Dev. |  | 04: 4.146%; 10: 5.27%; 13: 0%
PDM3
| Z-Score (Absolute) |  | 0 | 6.58 | 0 | 4.5 | 12.59 | 5.39 | 8.99 | 0 | 1.5 | 1.5 | 0 | 0 | 0 | 0 | 0 | 1.35 | 0 | 4.5
| Analytical results for Ba in OREAS 45f (Certified Value 158 ppm) | [31 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES |  |  |  |  | 0
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  |  | 0
| 1 | 1 | 161.6256 | 162 | 146 | 161 | 203.8448 | 164 | 155.5 | 161 | 179 | 165.8 | 150 | 148.7 | 155.1351 | 159 | 161.5 | 166.5 | 160 | 155 | 157 | 165 | 152.5 | 191 | 172 | 167.6667 | 143.361 | 156.3762 |  |  |  |  | 1
| 1 | 2 | 157.6512 | 159 | 146 | 164 | 184.5929 | 161 | 157 | 159 | 178 | 169.9 | 154 | 150.3 | 155.457 | 157 | 161.5 | 166 | 160 | 155.5 | 154 | 162 | 150 | 154 | 168 | 162.3333 | 143.152 | 158.7725 |  |  |  |  | #N/A
| 1 | 3 | 160.064 | 157 | 153 | 156 | 198.6784 | 169 | 157.5 | 158 | 178 | 160.6 | 151 | 146.1 | 159.1542 | 166 | 166 | 167.5 | 170 | 156 | 153 | 166 | 161 | 161 | 173 | 160.6667 | 147.749 | 161.6357 |  |  |  |  | 16
| 1 | 4 | 157.2192 | 153 | 155 | 161 | 188.2842 | 165 | 156 | 167 | 182 | 164.2 | 152 | 140.4 | 161.5669 | 158 | 165 | 168 | 170 | 156.5 | 152 | 162 | 162.5 | 160 | 166 | 162.6667 | 147.724 | 158.0698 |  |  |  |  | 157.9904
| 1 | 5 | 160.6912 | 158 | 146 | 156 | 206.0136 | 161 | 153 | 160 | 179 | 164.5 | 152 | 144.8 | 157.179 | 158 | 164 | 166.5 | 160 | 157 | 155 | 164 | 157 | 153 | 169 | 162.3333 | 146.537 | 163.0047 |  |  |  |  | 140
| 1 | 6 | 157.6576 | 153 | 145 | 162 | 189.417 | 149 | 153 | 161 | 177 | 168.9 | 156 | 148.6 | 153.0171 | 158 | 160.5 | 163.5 | 180 | 154 | 160 | 162 | 159 | 157 | 166 | 162.6667 | 146.516 | 159.2608
| Mean |  | 159.1515 | 157 | 148.5 | 160 | 195.1385 | 161.5 | 155.3333 | 161 | 178.8333 | 165.65 | 152.5 | 146.4833 | 156.9182 | 159.3333 | 163.0833 | 166.3333 | 166.6667 | 155.6667 | 155.1667 | 163.5 | 157 | 162.6667 | 169 | 163.0556 | 145.8398 | 159.52
| Median |  | 158.8608 | 157.5 | 146 | 161 | 194.0477 | 162.5 | 155.75 | 160.5 | 178.5 | 165.15 | 152 | 147.35 | 156.318 | 158 | 162.75 | 166.5 | 165 | 155.75 | 154.5 | 163 | 158 | 158.5 | 168.5 | 162.5 | 146.5265 | 159.0167
| Std Dev. |  | 1.873 | 3.5214 | 4.3243 | 3.2863 | 8.9165 | 6.8044 | 1.9408 | 3.1623 | 1.7224 | 3.3934 | 2.1679 | 3.5763 | 3.0728 | 3.3267 | 2.223 | 1.5706 | 8.165 | 1.0801 | 2.9269 | 1.7607 | 4.8888 | 14.2361 | 2.9665 | 2.3797 | 2.074 | 2.4179
| Rel.Std.Dev. |  | 1.177% | 2.243% | 2.912% | 2.054% | 4.569% | 4.213% | 1.249% | 1.964% | 0.9631% | 2.049% | 1.422% | 2.441% | 1.958% | 2.088% | 1.363% | 0.9442% | 4.899% | 0.6939% | 1.886% | 1.077% | 3.114% | 8.752% | 1.755% | 1.459% | 1.422% | 1.516%
| PDM3 |  | 0.7349% | -0.6269% | -6.007% | 1.272% | 23.51% | 2.221% | -1.682% | 1.905% | 13.19% | 4.848% | -3.475% | -7.283% | -0.6786% | 0.85% | 3.224% | 5.281% | 5.492% | -1.471% | -1.787% | 3.487% | -0.6269% | 2.96% | 6.969% | 3.206% | -7.691% | 0.9681%
| Z-Score (Absolute) |  | 0.06 | 0.44 | 1.93 | 0.08 | 6.23 | 0.35 | 0.73 | 0.26 | 3.38 | 1.07 | 1.23 | 2.28 | 0.46 | 0.03 | 0.62 | 1.19 | NA | 0.67 | 0.76 | 0.7 | 0.44 | 0.55 | 1.66 | 0.62 | 2.39 | 0
Results from laboratory 30 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 45f (Certified Value 0.98 ppm) | [31 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.63468 | 1 | 1 | 1.02 | 1.9165 | 1.1 | 0.96 | < 5 | 1 | 0.8 | 1.01 | 1.0338 | 0.973 | 1.1 | 0.67 | 0.95 | 0.9 | 0.97 | 1 | 1 | 0.97 | 1.08 | 1 | 0.96 | 0.59 | 0.98969 |  |  |  |  | 1
| 1 | 2 | 0.62896 | 0.9 | 1 | 1.05 | 1.6342 | 1.07 | 0.92 | < 5 | 1 | 0.8 | 1.04 | 1.0445 | 1.03 | 1 | 0.99 | 0.94 | 0.83 | 0.96 | 1 | 0.9 | 0.95 | 1.1 | 1 | 0.91333 | 0.61 | 0.92017 |  |  |  |  | #N/A
| 1 | 3 | 0.64316 | 0.9 | 1 | 1.01 | 1.9196 | 1.14 | 0.84 | < 5 | 1 | 0.8 | 1 | 1.0572 | 1.15 | 1 | 0.91 | 0.95 | 0.83 | 0.94 | 0.9 | 0.9 | 1.13 | 1.15 | 1 | 0.9 | 0.59 | 0.9965 |  |  |  |  | 16
| 1 | 4 | 0.65912 | 1.1 | 1.1 | 1.06 | 1.6673 | 1.2 | 0.93 | < 5 | 1 | 0.8 | 1 | 1.0115 | 1.01 | 1 | 0.8 | 0.96 | 0.94 | 0.93 | 1.1 | 0.9 | 1.18 | 1.13 | 1 | 0.90333 | 0.62 | 0.99114 |  |  |  |  | 0.98214
| 1 | 5 | 0.62072 | 1.1 | 1.1 | 1.02 | 1.8208 | 1.04 | 0.99 | < 5 | 0.9 | 0.8 | 0.99 | 1.0626 | 0.87 | 1 | 0.88 | 0.91 | 0.86 | 0.93 | 1 | 1 | 1.24 | 1.09 | 1 | 0.95 | 0.6 | 0.99534 |  |  |  |  | 141
| 1 | 6 | 0.62516 | 1 | 1.1 | 1.02 | 1.8108 | 0.97 | 0.98 | < 5 | 0.9 | 0.8 | 1 | 1.0158 | 0.974 | 1 | 0.75 | 0.93 | 0.91 | 0.92 | 0.9 | 1 | 1.07 | 1.14 | 1 | 0.91333 | 0.58 | 0.99768
| Mean |  | 0.6353 | 1 | 1.05 | 1.03 | 1.7949 | 1.0867 | 0.93667 |  | 0.96667 | 0.8 | 1.0067 | 1.0376 | 1.0012 | 1.0167 | 0.83333 | 0.94 | 0.87833 | 0.94167 | 0.98333 | 0.95 | 1.09 | 1.115 | 1 | 0.92333 | 0.59833 | 0.98175
| Median |  | 0.63182 | 1 | 1.05 | 1.02 | 1.8158 | 1.085 | 0.945 |  | 1 | 0.8 | 1 | 1.0392 | 0.992 | 1 | 0.84 | 0.945 | 0.88 | 0.935 | 1 | 0.95 | 1.1 | 1.115 | 1 | 0.91333 | 0.595 | 0.99324
| Std Dev. |  | 0.014037 | 0.089443 | 0.054772 | 0.02 | 0.12113 | 0.079917 | 0.05465 |  | 0.05164 | 1.2162e-16 | 0.017512 | 0.021101 | 0.09143 | 0.040825 | 0.11605 | 0.017889 | 0.04535 | 0.019408 | 0.075277 | 0.054772 | 0.11541 | 0.02881 | 0 | 0.025298 | 0.01472 | 0.030329
| Rel.Std.Dev. |  | 2.209% | 8.944% | 5.216% | 1.942% | 6.749% | 7.354% | 5.835% |  | 5.342% | 1.52e-14% | 1.74% | 2.034% | 9.132% | 4.016% | 13.93% | 1.903% | 5.163% | 2.061% | 7.655% | 5.766% | 10.59% | 2.584% | 0% | 2.74% | 2.46% | 3.089%
| PDM3 |  | -35.31% | 1.818% | 6.909% | 4.873% | 82.75% | 10.64% | -4.63% |  | -1.576% | -18.55% | 2.497% | 5.642% | 1.937% | 3.515% | -15.15% | -4.291% | -10.57% | -4.121% | 0.1213% | -3.273% | 10.98% | 13.53% | 1.818% | -5.988% | -39.08% | -0.03958%
| Z-Score (Absolute) |  | 4.51 | 0.11 | 0.74 | 0.48 | 10.16 | 1.2 | 0.7 | 19.07 | 0.32 | 2.42 | 0.19 | 0.58 | 0.12 | 0.32 | 2 | 0.65 | 1.43 | 0.63 | 0.11 | 0.53 | 1.24 | 1.56 | 0.11 | 0.86 | 4.97 | 0.13
| Analytical results for Bi in OREAS 45f (Certified Value 0.17 ppm) | [31 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | Various [a] | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 15g | see footer | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | < 0.5 | 0.16 | 0.16 | 0.17 | 0.23694 | 0.17 | 0.16 | 2 | 0.2 | 0.1 | 0.2 | 0.1 | 0.16 | 0.16464 | 0.159 | 0.33 | 0.152 | 0.15 | 0.16 | 0.16 | 0.2 | 0.18 | 0.147 | 0.2 | < 5 | 0.17333 | 0.079 | 0.15978 |  |  | 1
| 1 | 2 | < 0.5 | 0.15 | 0.17 | 0.18 | 0.30437 | 0.17 | 0.159 | 2 | 0.2 | 0.1 | 0.22 | 0.1 | 0.17 | 0.16127 | 0.162 | 0.34 | 0.166 | 0.151 | 0.15 | 0.162 | 0.1 | 0.18 | 0.154 | 0.19 | < 5 | 0.17333 | 0.067 | 0.15979 |  |  | #N/A
| 1 | 3 | < 0.5 | 0.17 | 0.16 | 0.17 | 0.25116 | 0.17 | 0.162 | 2 | 0.2 | 0.1 | 0.21 | < 5 | 0.16 | 0.15823 | 0.171 | 0.27 | 0.172 | 0.151 | 0.15 | 0.16 | 0.2 | 0.18 | 0.179 | 0.19 | < 5 | 0.17333 | 0.076 | 0.15885 |  |  | 16
| 1 | 4 | < 0.5 | 0.17 | 0.17 | 0.18 | 0.33748 | 0.18 | 0.164 | 3 | 0.2 | 0.1 | 0.2 | < 5 | 0.16 | 0.1651 | 0.171 | 0.28 | 0.158 | 0.152 | 0.15 | 0.168 | 0.1 | 0.17 | 0.176 | 0.2 | < 5 | 0.18 | 0.074 | 0.16119 |  |  | 0.16801
| 1 | 5 | < 0.5 | 0.16 | 0.17 | 0.17 | 0.29394 | 0.17 | 0.161 | 3 | 0.2 | 0.1 | 0.21 | 0.1 | 0.16 | 0.1598 | 0.164 | 0.3 | 0.17 | 0.154 | 0.15 | 0.163 | 0.1 | 0.18 | 0.168 | 0.2 | < 5 | 0.17333 | 0.072 | 0.15798 |  |  | 142
| 1 | 6 | < 0.5 | 0.16 | 0.16 | 0.18 | 0.32217 | 0.16 | 0.156 | 2 | 0.2 | 0.1 | 0.2 | 0.1 | 0.16 | 0.16572 | 0.164 | 0.3 | 0.158 | 0.152 | 0.15 | 0.157 | 0.1 | 0.18 | 0.158 | 0.2 | < 5 | 0.17333 | 0.073 | 0.15747
| Mean |  |  | 0.16167 | 0.165 | 0.175 | 0.29101 | 0.17 | 0.16033 | 2.3333 | 0.2 | 0.1 | 0.20667 | 0.1 | 0.16167 | 0.16246 | 0.16517 | 0.30333 | 0.16267 | 0.15167 | 0.15167 | 0.16167 | 0.13333 | 0.17833 | 0.16367 | 0.19667 |  | 0.17444 | 0.0735 | 0.15918
| Median |  |  | 0.16 | 0.165 | 0.175 | 0.29916 | 0.17 | 0.1605 | 2 | 0.2 | 0.1 | 0.205 | 0.1 | 0.16 | 0.16296 | 0.164 | 0.3 | 0.162 | 0.1515 | 0.15 | 0.161 | 0.1 | 0.18 | 0.163 | 0.2 |  | 0.17333 | 0.0735 | 0.15932
| Std Dev. |  |  | 0.0075277 | 0.0054772 | 0.0054772 | 0.039569 | 0.0063246 | 0.0027325 | 0.5164 | 3.0405e-17 | 1.5202e-17 | 0.008165 | 0 | 0.0040825 | 0.0031207 | 0.0048751 | 0.027325 | 0.0078655 | 0.0013663 | 0.0040825 | 0.0037238 | 0.05164 | 0.0040825 | 0.012723 | 0.005164 |  | 0.0027217 | 0.0040373 | 0.0013598
| Rel.Std.Dev. |  |  | 4.656% | 3.32% | 3.13% | 13.6% | 3.72% | 1.704% | 22.13% | 1.52e-14% | 1.52e-14% | 3.951% | 0% | 2.525% | 1.921% | 2.952% | 9.008% | 4.835% | 0.9008% | 2.692% | 2.303% | 38.73% | 2.289% | 7.774% | 2.626% |  | 1.56% | 5.493% | 0.8543%
| PDM3 |  |  | -3.774% | -1.79% | 4.162% | 73.21% | 1.186% | -4.568% | 1289% | 19.04% | -40.48% | 23.01% | -40.48% | -3.774% | -3.302% | -1.691% | 80.55% | -3.179% | -9.726% | -9.726% | -3.774% | -20.64% | 6.146% | -2.584% | 17.06% |  | 3.831% | -56.25% | -5.256%
| Z-Score (Absolute) |  | 5.56 | 0.17 | 0.04 | 0.69 | 8.22 | 0.37 | 0.26 | NA | NA | 4.18 | 2.75 | NA | 0.17 | 0.12 | 0.05 | 9.02 | 0.11 | 0.82 | 0.82 | 0.17 | NA | 0.91 | 0.04 | 2.1 | 151.6 | 0.66 | 5.9 | 0.33
Results from laboratory 17, 18, 21 & 33 were removed due to their 0.1 ppm or greater reading resolution.
[a]: Lab 21: Rnd 01 - AR*MS (0.5g), Rnd 01 - AR*OES (0.15g)
| Analytical results for Ca in OREAS 45f (Certified Value 0.075 wt.%) | [31 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.070485 | 0.08 | 0.07 | 0.06 | 0.085614 | 0.08 | 0.07 | 0.0811 | 0.08 | 0.076 | 0.08 | 0.07 | 0.07 | 0.072 | 0.061779 | 0.0744 | 0.08 | 0.07 | 0.07 | 0.07 | 0.08 | 0.0705 | 0.07 | 0.07 | 0.1 | 0.08 | 0.08 | 0.0784 | 0.077282 |  | 1
| 1 | 2 | 0.070921 | 0.07 | 0.07 | 0.06 | 0.084747 | 0.08 | 0.07 | 0.0776 | 0.07 | 0.0755 | 0.08 | 0.07 | 0.07 | 0.075 | 0.061645 | 0.0727 | 0.08 | 0.08 | 0.07 | 0.08 | 0.08 | 0.065 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.0791 | 0.079019 |  | #N/A
| 1 | 3 | 0.070573 | 0.08 | 0.07 | 0.06 | 0.08618 | 0.08 | 0.07 | 0.0838 | 0.08 | 0.0755 | 0.08 | 0.08 | 0.07 | 0.073 | 0.063565 | 0.0746 | 0.08 | 0.08 | 0.07 | 0.08 | 0.07 | 0.064 | 0.07 | 0.08 | 0.08 | 0.1 | 0.083333 | 0.079 | 0.079956 |  | 16
| 1 | 4 | 0.071084 | 0.07 | 0.07 | 0.06 | 0.085562 | 0.08 | 0.07 | 0.0804 | 0.08 | 0.0785 | 0.08 | 0.07 | 0.07 | 0.069 | 0.068595 | 0.0746 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.0665 | 0.07 | 0.08 | 0.08 | 0.08 | 0.083333 | 0.0783 | 0.080633 |  | 0.075053
| 1 | 5 | 0.070794 | 0.08 | 0.07 | 0.06 | 0.086197 | 0.08 | 0.07 | 0.0786 | 0.08 | 0.0765 | 0.08 | 0.08 | 0.07 | 0.069 | 0.066871 | 0.074 | 0.08 | 0.08 | 0.07 | 0.07 | 0.08 | 0.0685 | 0.07 | 0.08 | 0.07 | 0.09 | 0.083333 | 0.0783 | 0.081665 |  | 143
| 1 | 6 | 0.070973 | 0.07 | 0.07 | 0.06 | 0.085469 | 0.08 | 0.07 | 0.0802 | 0.08 | 0.077 | 0.08 | 0.07 | 0.07 | 0.073 | 0.063688 | 0.0718 | 0.09 | 0.07 | 0.07 | 0.08 | 0.08 | 0.0715 | 0.07 | 0.08 | 0.08 | 0.08 | 0.083333 | 0.0788 | 0.080978
| Mean |  | 0.070805 | 0.075 | 0.07 | 0.06 | 0.085628 | 0.08 | 0.07 | 0.080283 | 0.078333 | 0.0765 | 0.08 | 0.073333 | 0.07 | 0.071833 | 0.064357 | 0.073683 | 0.081667 | 0.076667 | 0.07 | 0.075 | 0.078333 | 0.067667 | 0.07 | 0.076667 | 0.081667 | 0.085 | 0.082222 | 0.07865 | 0.079922
| Median |  | 0.070858 | 0.075 | 0.07 | 0.06 | 0.085588 | 0.08 | 0.07 | 0.0803 | 0.08 | 0.07625 | 0.08 | 0.07 | 0.07 | 0.0725 | 0.063626 | 0.0742 | 0.08 | 0.08 | 0.07 | 0.075 | 0.08 | 0.0675 | 0.07 | 0.08 | 0.08 | 0.08 | 0.083333 | 0.0786 | 0.080295
| Std Dev. |  | 0.00023489 | 0.0054772 | 0 | 0 | 0.0005358 | 0 | 0 | 0.0021489 | 0.0040825 | 0.0011402 | 0 | 0.005164 | 0 | 0.0024014 | 0.0028056 | 0.0011669 | 0.0040825 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.0030111 | 0 | 0.005164 | 0.0098319 | 0.0083666 | 0.0017213 | 0.00036194 | 0.0015776
| Rel.Std.Dev. |  | 0.3317% | 7.303% | 0% | 0% | 0.6257% | 0% | 0% | 2.677% | 5.212% | 1.49% | 0% | 7.042% | 0% | 3.343% | 4.359% | 1.584% | 4.999% | 6.736% | 0% | 7.303% | 5.212% | 4.45% | 0% | 6.736% | 12.04% | 9.843% | 2.094% | 0.4602% | 1.974%
| PDM3 |  | -5.661% | -0.07127% | -6.733% | -20.06% | 14.09% | 6.591% | -6.733% | 6.968% | 4.37% | 1.927% | 6.591% | -2.292% | -6.733% | -4.29% | -14.25% | -1.826% | 8.811% | 2.149% | -6.733% | -0.07127% | 4.37% | -9.842% | -6.733% | 2.149% | 8.811% | 13.25% | 9.551% | 4.792% | 6.487%
| Z-Score (Absolute) |  | 0.82 | 0.22 | 0.94 | 2.38 | 1.32 | 0.51 | 0.94 | 0.55 | 0.26 | 0 | 0.51 | 0.46 | 0.94 | 0.67 | 1.75 | 0.41 | 0.75 | 0.02 | 0.94 | 0.22 | 0.26 | 1.28 | 0.94 | 0.02 | 0.75 | 1.23 | 0.83 | 0.31 | 0.49
| Analytical results for Cd in OREAS 45f (Indicative Value 0.011 ppm) | [31 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  |  |  | 3
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02 | < 0.01 | 0.01 | < 0.1 | 0.01 | 0.006 | 0.24 | < 0.5 | < 0.1 | < 0.1 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | 0.0112 | 0.009 | 0.006 | 0.01 | 0.005 | < 0.1 | < 0.01 | 0.008 | 0.01 | < 1 | 0.02 | 0.105 | 0.00631 |  |  |  | 1
| 1 | 2 | 0.02 | < 0.01 | 0.01 | < 0.1 | 0.01 | 0.007 | 0.21 | < 0.5 | < 0.1 | < 0.1 | < 0.01 | < 0.1 | 0.01 | < 0.01 | < 0.01 | 0.007 | < 0.001 | < 0.01 | 0.004 | < 0.1 | 0.01 | 0.008 | 0.02 | < 1 | 0.013333 | 0.12 | 0.00909 |  |  |  | 10
| 1 | 3 | 0.03 | < 0.01 | 0.01 | < 0.1 | < 0.01 | 0.006 | 0.28 | < 0.5 | < 0.1 | < 0.1 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | 0.0163 | 0.011 | 0.004 | 0.01 | 0.011 | < 0.1 | 0.01 | 0.003 | < 0.01 | < 1 | 0.02 | 0.11 | 0.00532 |  |  |  | 16
| 1 | 4 | 0.02 | < 0.01 | 0.02 | < 0.1 | 0.02 | 0.001 | 0.26 | < 0.5 | < 0.1 | < 0.1 | < 0.01 | < 0.1 | < 0.01 | < 0.01 | < 0.01 | < 0.001 | 0.001 | 0.01 | 0.005 | < 0.1 | < 0.01 | 0.004 | < 0.01 | < 1 | 0.02 | 0.11 | 0.01105 |  |  |  | 0.011154
| 1 | 5 | 0.03 | < 0.01 | 0.01 | < 0.1 | 0.01 | 0.003 | 0.23 | < 0.5 | < 0.1 | < 0.1 | < 0.01 | < 0.1 | 0.01 | < 0.01 | < 0.01 | 0.009 | 0.009 | < 0.01 | 0.005 | < 0.1 | < 0.01 | 0.008 | < 0.01 | < 1 | 0.02 | 0.11 | 0.00736 |  |  |  | 30
| 1 | 6 | 0.03 | < 0.01 | 0.01 | < 0.1 | < 0.01 | 0.006 | 0.24 | < 0.5 | < 0.1 | < 0.1 | < 0.01 | 0.1 | < 0.01 | < 0.01 | 0.0116 | < 0.001 | 0.002 | 0.01 | 0.005 | < 0.1 | < 0.01 | 0.006 | 0.02 | < 1 | 0.023333 | 0.1 | 0.00776
| Mean |  | 0.025 |  | 0.011667 |  | 0.0125 | 0.0048333 | 0.24333 |  |  |  |  | 0.1 | 0.01 |  | 0.013033 | 0.009 | 0.0044 | 0.01 | 0.0058333 |  | 0.01 | 0.0061667 | 0.016667 |  | 0.019444 | 0.10917 | 0.007815
| Median |  | 0.025 |  | 0.01 |  | 0.01 | 0.006 | 0.24 |  |  |  |  | 0.1 | 0.01 |  | 0.0116 | 0.009 | 0.004 | 0.01 | 0.005 |  | 0.01 | 0.007 | 0.02 |  | 0.02 | 0.11 | 0.00756
| Std Dev. |  | 0.0054772 |  | 0.0040825 |  | 0.005 | 0.0023166 | 0.024221 | [5 empty] | 0 |  | 0.0028361 | 0.001633 | 0.0032094 | 0 | 0.0025626 |  | 0 | 0.0022286 | 0.0057735 |  | 0.0032773 | 0.0066458 | 0.0020392
| Rel.Std.Dev. |  | 21.91% |  | 34.99% |  | 40% | 47.93% | 9.954% | [5 empty] | 0% |  | 21.76% | 18.14% | 72.94% | 0% | 43.93% |  | 0% | 36.14% | 34.64% |  | 16.85% | 6.088% | 26.09%
| PDM3 |  | 124.1% |  | 4.593% |  | 12.06% | -56.67% | 2082% |  |  |  |  | 796.5% | -10.35% |  | 16.85% | -19.31% | -60.55% | -10.35% | -47.7% |  | -10.35% | -44.72% | 49.42% |  | 74.32% | 878.7% | -29.94%
| Z-Score (Absolute) |  | 2.02 | 0.67 | 0.22 | 5.39 | 0 | 0.7 | 31.47 | 32.37 | 5.39 | 5.39 | 0.67 | 6.52 | 0.45 | 0.67 | 0.13 | 0.52 | 0.84 | 0.22 | 0.56 | 5.39 | 0.45 | 0.52 | 0.11 | 66.08 | 1.27 | 13.37 | 0.29
| Analytical results for Ce in OREAS 45f (Certified Value 22.3 ppm) | [31 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 18 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.4g | 0.15g | 0.5g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 25.0645 | 21.01 | 20.5 | 24.5 | 27.204 | 22.53 | 23.4 | 19 | 19.8 | 26 | 19.93 | 22.9623 | 20.6 | 22.3221 | 23.2 | 21.7 | 22 | 21.7 | 21.4 | 25.3 | 19.9 | 19.2 | 26.4 | 26.39 | 13.437 | 21.1527 |  |  |  |  | 1
| 1 | 2 | 24.645 | 20.76 | 21.2 | 24.5 | 24.3753 | 22.36 | 23 | 18 | 19.7 | 26 | 21.09 | 23.2131 | 20.1 | 20.5739 | 22.3 | 22.8 | 21.9 | 20.6 | 20.9 | 20.4 | 19.84 | 19 | 27.1 | 25.0367 | 13.867 | 21.4697 |  |  |  |  | #N/A
| 1 | 3 | 24.7568 | 20.28 | 21.7 | 23.8 | 26.3013 | 22.63 | 23.2 | 19 | 19.9 | 26 | 20.21 | 23.9491 | 21.2 | 20.6558 | 21.8 | 22.3 | 22.1 | 20.3 | 21 | 21 | 17.49 | 20.6 | 25.6 | 25.82 | 13.511 | 21.3434 |  |  |  |  | 16
| 1 | 4 | 24.8979 | 20.33 | 21.8 | 24.5 | 24.8752 | 22.92 | 23.4 | 19 | 20.2 | 27 | 19.88 | 24.7915 | 21.3 | 21.5113 | 21.8 | 21.8 | 22 | 21.3 | 21.1 | 21.8 | 19.1 | 21.3 | 25.4 | 26.0733 | 13.85 | 21.9612 |  |  |  |  | 22.3091
| 1 | 5 | 24.453 | 20.04 | 21.7 | 24.7 | 26.7497 | 22.09 | 22.8 | 19 | 20.3 | 27 | 20.38 | 24.7833 | 20.8 | 20.3451 | 21 | 22 | 22.2 | 20.5 | 20.9 | 23.2 | 19.42 | 20.4 | 25.9 | 25.64 | 13.973 | 21.7366 |  |  |  |  | 144
| 1 | 6 | 24.3207 | 20.12 | 21.3 | 24.5 | 25.1421 | 19.1 | 22.6 | 19 | 18.7 | 28 | 20.3 | 22.9645 | 20.4 | 23.2151 | 22.2 | 21.9 | 22 | 21.1 | 20.8 | 23.4 | 19.34 | 20.7 | 26.6 | 25.4533 | 14.001 | 21.4515
| Mean |  | 24.6896 | 20.4233 | 21.3667 | 24.4167 | 25.7746 | 21.9383 | 23.0667 | 18.8333 | 19.7667 | 26.6667 | 20.2983 | 23.7773 | 20.7333 | 21.4372 | 22.05 | 22.0833 | 22.0333 | 20.9167 | 21.0167 | 22.5167 | 19.1817 | 20.2 | 26.1667 | 25.7356 | 13.7732 | 21.5192
| Median |  | 24.7009 | 20.305 | 21.5 | 24.5 | 25.7217 | 22.445 | 23.1 | 19 | 19.85 | 26.5 | 20.255 | 23.5811 | 20.7 | 21.0836 | 22 | 21.95 | 22 | 20.85 | 20.95 | 22.5 | 19.38 | 20.5 | 26.15 | 25.73 | 13.8585 | 21.4606
| Std Dev. |  | 0.2766 | 0.38098 | 0.48854 | 0.31252 | 1.1348 | 1.4177 | 0.3266 | 0.40825 | 0.57155 | 0.8165 | 0.43632 | 0.86178 | 0.46332 | 1.14 | 0.72595 | 0.40702 | 0.10328 | 0.53821 | 0.2137 | 1.8049 | 0.88314 | 0.90554 | 0.64704 | 0.47471 | 0.24012 | 0.28805
| Rel.Std.Dev. |  | 1.12% | 1.865% | 2.286% | 1.28% | 4.403% | 6.462% | 1.416% | 2.168% | 2.891% | 3.062% | 2.15% | 3.624% | 2.235% | 5.318% | 3.292% | 1.843% | 0.4687% | 2.573% | 1.017% | 8.016% | 4.604% | 4.483% | 2.473% | 1.845% | 1.743% | 1.339%
| PDM3 |  | 10.67% | -8.453% | -4.224% | 9.447% | 15.53% | -1.662% | 3.396% | -15.58% | -11.4% | 19.53% | -9.013% | 6.581% | -7.063% | -3.908% | -1.161% | -1.012% | -1.236% | -6.242% | -5.793% | 0.9304% | -14.02% | -9.454% | 17.29% | 15.36% | -38.26% | -3.541%
| Z-Score (Absolute) |  | 1.44 | 0.64 | 0.18 | 1.31 | 1.97 | 0.1 | 0.65 | 1.41 | 0.96 | 2.41 | 0.7 | 1 | 0.48 | 0.14 | 0.16 | 0.17 | 0.15 | 0.4 | 0.35 | 0.38 | 1.24 | 0.74 | 2.16 | 1.95 | 3.88 | 0.1
| Analytical results for Co in OREAS 45f (Certified Value 39.2 ppm) | [31 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 40.726 | 40 | 40.1 | 44.1 | 41.1286 | 43.5 | 39 | 43 | 36 | 32.7 | 38.2 | 39.7 | 34.5 | 38.66 | 42.6392 | 35.6 | 38.9675 | 39.4 | 40.6 | 37 | 37.7 | 41.7 | 35 | 42.1 | 37.5 | 45 | 33 | 47.6 | 17.204 | 38.525 | 1
| 1 | 2 | 40.1388 | 41 | 39.9 | 46.3 | 39.891 | 42 | 39.1 | 41 | 36 | 32.2 | 38.2 | 40.2 | 33.9 | 39.55 | 43.144 | 36.5 | 38.3742 | 38.4 | 38 | 37.3 | 37.1 | 40.8 | 35 | 41.2 | 38 | 45.1 | 33 | 45 | 16.883 | 39.1639 | 3
| 1 | 3 | 40.3517 | 38 | 40.8 | 43.6 | 41.7131 | 42.8 | 39.1 | 43 | 36 | 32.8 | 37.2 | 39.3 | 33.7 | 38.41 | 43.9751 | 35.2 | 37.8439 | 37.7 | 42.5 | 37.8 | 36.1 | 41.2 | 34 | 38.9 | 40.5 | 44.1 | 33 | 45.4333 | 18.159 | 39.4034 | 16
| 1 | 4 | 41.0538 | 39 | 41.6 | 45.1 | 39.3445 | 44 | 38.5 | 42 | 36 | 32.8 | 37.6 | 40.3 | 34 | 36.47 | 43.9488 | 36.6 | 38.8633 | 39.4 | 41.9 | 38 | 38.3 | 42.2 | 34 | 40.7 | 40.3 | 44.3 | 32 | 46.1 | 17.645 | 38.8318 | 39.2404
| 1 | 5 | 40.3589 | 39 | 40.1 | 42.4 | 41.4201 | 42.3 | 37.3 | 42 | 35 | 32.5 | 37.4 | 40.1 | 34.1 | 37.37 | 42.5959 | 35.4 | 39.2644 | 38 | 42.3 | 37 | 36.9 | 40.9 | 35 | 41 | 39.6 | 45.8 | 33 | 46.3667 | 18.1 | 39.9129 | 145
| 1 | 6 | 40.6353 | 39 | 39.9 | 45.5 | 40.9574 | 37.9 | 36.8 | 43 | 35 | 30.3 | 38.4 | 38.7 | 33.9 | 38.59 | 43.1869 | 35.6 | 38.8797 | 40.9 | 40.9 | 36.7 | 38.4 | 40.8 | 35 | 41.8 | 40.5 | 42.5 | 32 | 45.9 | 18.138 | 39.1673
| Mean |  | 40.5441 | 39.3333 | 40.4 | 44.5 | 40.7425 | 42.0833 | 38.3 | 42.3333 | 35.6667 | 32.2167 | 37.8333 | 39.7167 | 34.0167 | 38.175 | 43.2483 | 35.8167 | 38.6988 | 38.9667 | 41.0333 | 37.3 | 37.4167 | 41.2667 | 34.6667 | 40.95 | 39.4 | 44.4667 | 32.6667 | 46.0667 | 17.6882 | 39.1674
| Median |  | 40.4971 | 39 | 40.1 | 44.6 | 41.043 | 42.55 | 38.75 | 42.5 | 36 | 32.6 | 37.9 | 39.9 | 33.95 | 38.5 | 43.1654 | 35.6 | 38.8715 | 38.9 | 41.4 | 37.15 | 37.4 | 41.05 | 35 | 41.1 | 39.95 | 44.65 | 33 | 46 | 17.8725 | 39.1656
| Std Dev. |  | 0.32766 | 1.0328 | 0.67528 | 1.4128 | 0.92477 | 2.1794 | 1.006 | 0.8165 | 0.5164 | 0.96626 | 0.49666 | 0.62102 | 0.27142 | 1.087 | 0.60502 | 0.58793 | 0.50754 | 1.181 | 1.6681 | 0.50596 | 0.88638 | 0.57155 | 0.5164 | 1.1292 | 1.3297 | 1.1396 | 0.5164 | 0.89542 | 0.54373 | 0.47681
| Rel.Std.Dev. |  | 0.8082% | 2.626% | 1.671% | 3.175% | 2.27% | 5.179% | 2.627% | 1.929% | 1.448% | 2.999% | 1.313% | 1.564% | 0.7979% | 2.848% | 1.399% | 1.642% | 1.312% | 3.031% | 4.065% | 1.356% | 2.369% | 1.385% | 1.49% | 2.757% | 3.375% | 2.563% | 1.581% | 1.944% | 3.074% | 1.217%
| PDM3 |  | 3.322% | 0.2367% | 2.955% | 13.4% | 3.828% | 7.245% | -2.397% | 7.882% | -9.107% | -17.9% | -3.586% | 1.214% | -13.31% | -2.715% | 10.21% | -8.725% | -1.38% | -0.6977% | 4.569% | -4.945% | -4.648% | 5.164% | -11.66% | 4.357% | 0.4066% | 13.32% | -16.75% | 17.4% | -54.92% | -0.1862%
| Z-Score (Absolute) |  | 0.46 | 0.03 | 0.41 | 1.87 | 0.53 | 1.01 | 0.34 | 1.1 | 1.28 | 2.51 | 0.51 | 0.17 | 1.87 | 0.38 | 1.42 | 1.22 | 0.2 | 0.1 | 0.64 | 0.7 | 0.65 | 0.72 | 1.63 | 0.61 | 0.05 | 1.86 | 2.35 | 2.43 | 7.68 | 0.03
| Analytical results for Cr in OREAS 45f (Certified Value 341 ppm) | [31 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 364.81 | 356 | 334 | 367 | 314.64 | 344 | 336 | 381 | 350 | 302 | 310 | 319 | 312 | 344 | 343.787 | 351 | 315 | 348 | 347 | 343 | 339 | 353 | 330 | 332 | 301 | 315 | 398.6667 | 172.653 | 379.9152 |  | 1
| 1 | 2 | 365.11 | 342 | 335 | 388 | 281.61 | 331 | 333 | 361 | 350 | 298 | 300 | 327.6 | 307 | 354.9 | 369.9269 | 347 | 307 | 355 | 345 | 361 | 342 | 334 | 327 | 322 | 364 | 313 | 392.6667 | 169 | 370.1453 |  | #N/A
| 1 | 3 | 362.2 | 344 | 347 | 378 | 278.2 | 344 | 339 | 371 | 350 | 294 | 320 | 305.1 | 312 | 339.2 | 356.5414 | 368 | 300 | 352 | 347 | 363 | 345 | 328 | 317 | 338 | 381 | 323 | 396.3333 | 178.867 | 381.8344 |  | 16
| 1 | 4 | 365.31 | 336 | 363 | 378 | 303.85 | 332 | 333 | 371 | 350 | 306 | 340 | 317.6 | 310 | 330.2 | 352.3832 | 350 | 314 | 350 | 350 | 352 | 344 | 335 | 325 | 340 | 361 | 312 | 399.6667 | 174.287 | 371.6924 |  | 341.1928
| 1 | 5 | 362.3 | 343 | 340 | 389 | 324.98 | 327 | 329 | 365 | 350 | 294 | 310 | 319.3 | 309 | 332 | 367.6424 | 356 | 309 | 348 | 344 | 349 | 344 | 344 | 321 | 330 | 355 | 317 | 395 | 178.856 | 379.1891 |  | 146
| 1 | 6 | 365.05 | 348 | 333 | 383 | 297.74 | 333 | 328 | 376 | 340 | 296 | 310 | 303.8 | 307 | 342.9 | 351.3906 | 349 | 330 | 340 | 345 | 365 | 336 | 361 | 328 | 335 | 364 | 314 | 393.3333 | 176.765 | 372.4748
| Mean |  | 364.13 | 344.8333 | 342 | 380.5 | 300.17 | 335.1667 | 333 | 370.8333 | 348.3333 | 298.3333 | 315 | 315.4 | 309.5 | 340.5333 | 356.9453 | 353.5 | 312.5 | 348.8333 | 346.3333 | 355.5 | 341.6667 | 342.5 | 324.6667 | 332.8333 | 354.3333 | 315.6667 | 395.9444 | 175.0713 | 375.8752
| Median |  | 364.93 | 343.5 | 337.5 | 380.5 | 300.795 | 332.5 | 333 | 371 | 350 | 297 | 310 | 318.3 | 309.5 | 341.05 | 354.4623 | 350.5 | 311.5 | 349 | 346 | 356.5 | 343 | 339.5 | 326 | 333.5 | 362.5 | 314.5 | 395.6667 | 175.526 | 375.832
| Std Dev. |  | 1.4653 | 6.7057 | 11.5239 | 8.1179 | 18.284 | 7.1391 | 4.1473 | 7.2226 | 4.0825 | 4.8028 | 13.784 | 9.1924 | 2.2583 | 9.0044 | 10.0789 | 7.7136 | 10.1341 | 5.0761 | 2.1602 | 8.8034 | 3.5024 | 12.5976 | 4.8442 | 6.4627 | 27.5221 | 3.9833 | 2.8238 | 3.871 | 4.9941
| Rel.Std.Dev. |  | 0.4024% | 1.945% | 3.37% | 2.133% | 6.091% | 2.13% | 1.245% | 1.948% | 1.172% | 1.61% | 4.376% | 2.915% | 0.7297% | 2.644% | 2.824% | 2.182% | 3.243% | 1.455% | 0.6237% | 2.476% | 1.025% | 3.678% | 1.492% | 1.942% | 7.767% | 1.262% | 0.7132% | 2.211% | 1.329%
| PDM3 |  | 6.723% | 1.067% | 0.2366% | 11.52% | -12.02% | -1.766% | -2.401% | 8.687% | 2.093% | -12.56% | -7.677% | -7.56% | -9.289% | -0.1933% | 4.617% | 3.607% | -8.41% | 2.239% | 1.507% | 4.193% | 0.1389% | 0.3831% | -4.844% | -2.45% | 3.851% | -7.481% | 16.05% | -48.69% | 10.17%
| Z-Score (Absolute) |  | 1 | 0.13 | 0 | 1.74 | 1.89 | 0.31 | 0.41 | 1.3 | 0.29 | 1.97 | 1.22 | 1.2 | 1.47 | 0.07 | 0.67 | 0.52 | 1.33 | 0.31 | 0.2 | 0.61 | 0.02 | 0.02 | 0.78 | 0.41 | 0.56 | 1.19 | 2.43 | 7.53 | 1.53
| Analytical results for Cs in OREAS 45f (Certified Value 1.88 ppm) | [31 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 18 | 19 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.4g | 40g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 3.0766 | 1.97 | 1.97 | 2.85 | 2.3169 | 2.8 | 1.95 | 1.31 | 1.5 | 1.81 | 3.0423 | 1.54 | 2.09 | 2.08 | 1.82 | 1.83 | 1.96 | 2.37 | 1.89 | 1.82 | 2.23 | 3.4267 | 0.598 | 1.8671 | [6 empty] | 1
| 1 | 2 | 3.0218 | 2.06 | 2.1 | 2.96 | 2.0305 | 2.69 | 1.98 | 1.28 | 1.4 | 1.8 | 3.0754 | 1.58 | 1.77 | 2.15 | 1.85 | 1.78 | 1.96 | 1.98 | 1.8 | 1.835 | 2.32 | 3.2833 | 0.566 | 1.8025 | [6 empty] | #N/A
| 1 | 3 | 3.0136 | 2.04 | 2.17 | 2.8 | 2.2371 | 2.61 | 2.01 | 1.26 | 1.5 | 1.75 | 3.1781 | 1.85 | 1.69 | 2.22 | 1.925 | 1.7 | 1.91 | 2.04 | 1.61 | 2.04 | 2.16 | 3.1833 | 0.627 | 1.912 | [6 empty] | 16
| 1 | 4 | 3.0711 | 1.95 | 2.15 | 2.9 | 2.0862 | 2.99 | 1.96 | 1.3 | 1.3 | 1.75 | 3.197 | 1.72 | 1.61 | 2.11 | 1.86 | 1.75 | 1.935 | 1.93 | 1.84 | 2.11 | 2.11 | 3.3267 | 0.573 | 1.9373 | [6 empty] | 1.876
| 1 | 5 | 2.9696 | 2.05 | 2.16 | 2.77 | 2.2907 | 2.46 | 1.89 | 1.19 | 1.2 | 1.72 | 3.1982 | 1.85 | 1.77 | 2.1 | 1.815 | 1.68 | 1.88 | 2.17 | 1.87 | 1.91 | 2.24 | 3.24 | 0.593 | 1.9677 | [6 empty] | 147
| 1 | 6 | 2.987 | 1.97 | 2.12 | 2.86 | 2.2117 | 1.69 | 1.875 | 1.2 | 1.4 | 1.72 | 3.1063 | 1.62 | 1.8 | 2.07 | 1.85 | 1.79 | 1.835 | 2.13 | 1.86 | 1.93 | 2.27 | 3.24 | 0.61 | 1.9227
| Mean |  | 3.0233 | 2.0067 | 2.1117 | 2.8567 | 2.1955 | 2.54 | 1.9442 | 1.2567 | 1.3833 | 1.7583 | 3.1329 | 1.6933 | 1.7883 | 2.1217 | 1.8533 | 1.755 | 1.9133 | 2.1033 | 1.8117 | 1.9408 | 2.2217 | 3.2833 | 0.5945 | 1.9015
| Median |  | 3.0177 | 2.005 | 2.135 | 2.855 | 2.2244 | 2.65 | 1.955 | 1.27 | 1.4 | 1.75 | 3.1422 | 1.67 | 1.77 | 2.105 | 1.85 | 1.765 | 1.9225 | 2.085 | 1.85 | 1.92 | 2.235 | 3.2617 | 0.5955 | 1.9173
| Std Dev. |  | 0.043413 | 0.048442 | 0.07414 | 0.068313 | 0.11396 | 0.45308 | 0.052194 | 0.050859 | 0.1169 | 0.038687 | 0.06728 | 0.1353 | 0.16327 | 0.055648 | 0.039455 | 0.056833 | 0.04916 | 0.15845 | 0.10342 | 0.1143 | 0.075741 | 0.085062 | 0.02274 | 0.058684
| Rel.Std.Dev. |  | 1.436% | 2.414% | 3.511% | 2.391% | 5.19% | 17.84% | 2.685% | 4.047% | 8.451% | 2.2% | 2.148% | 7.99% | 9.13% | 2.623% | 2.129% | 3.238% | 2.569% | 7.533% | 5.709% | 5.889% | 3.409% | 2.591% | 3.825% | 3.086%
| PDM3 |  | 61.15% | 6.963% | 12.56% | 52.27% | 17.03% | 35.39% | 3.631% | -33.01% | -26.26% | -6.274% | 66.99% | -9.739% | -4.675% | 13.09% | -1.21% | -6.452% | 1.988% | 12.12% | -3.431% | 3.454% | 18.42% | 75.01% | -68.31% | 1.359%
| Z-Score (Absolute) |  | 3.92 | 0.23 | 0.61 | 3.32 | 0.92 | 2.17 | 0.01 | 2.49 | 2.03 | 0.67 | 4.32 | 0.9 | 0.56 | 0.65 | 0.32 | 0.68 | 0.11 | 0.58 | 0.47 | 0.01 | 1.01 | 4.87 | 4.89 | 0.15
| Analytical results for Cu in OREAS 45f (Certified Value 336 ppm) | [31 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 355.58 | 329.1 | 334 | 343 | 372.7877 | 338.7 | 324 | 341 | 321 | 284 | 328 | 337.46 | 320.9 | 332.3 | 351.0715 | 335 | 365.7865 | 326 | 333 | 334 | 330 | 316 | 348.6 | 309 | 399 | 320 | 367.0333 | 128 | 334.4607 |  | 1
| 1 | 2 | 356.56 | 340.7 | 332 | 352 | 358.4736 | 329.7 | 328 | 331 | 322 | 279 | 320 | 338.56 | 318.3 | 322.9 | 351.8161 | 331 | 345.9333 | 337 | 334 | 350 | 334 | 325 | 331.7 | 306 | 342 | 317 | 356.0333 | 118 | 333.8129 |  | #N/A
| 1 | 3 | 355.57 | 342.2 | 358 | 340 | 357.2956 | 339.6 | 327 | 338 | 315 | 273 | 328 | 333.66 | 315.1 | 315.3 | 359.9068 | 336 | 343.5442 | 333 | 329 | 354 | 335 | 326 | 320.1 | 323 | 359 | 322 | 367.2333 | 130.255 | 332.827 |  | 16
| 1 | 4 | 356.79 | 331.2 | 352 | 353 | 346.6426 | 341.2 | 324 | 340 | 329 | 285 | 320 | 338.72 | 313.5 | 341.8 | 358.6322 | 339 | 350.2605 | 324 | 339 | 343 | 333 | 316 | 336.1 | 328 | 356 | 309 | 365.2 | 120.743 | 317.9765 |  | 335.8292
| 1 | 5 | 356.89 | 344.7 | 334 | 351 | 379.3216 | 329.8 | 314 | 330 | 312 | 271 | 306 | 346.69 | 316.5 | 310.9 | 354.3814 | 333 | 359.4547 | 327 | 329 | 341 | 337 | 311 | 341.1 | 314 | 342 | 313 | 366.4 | 127.149 | 337.266 |  | 148
| 1 | 6 | 359.58 | 338.9 | 326 | 350 | 364.2735 | 314.2 | 320 | 348 | 318 | 272 | 320 | 336.64 | 312.9 | 301.1 | 352.5369 | 324 | 371.5661 | 320 | 328 | 356 | 330 | 317 | 346 | 319 | 348 | 309 | 361.2333 | 130.319 | 323.9649
| Mean |  | 356.8283 | 337.8 | 339.3333 | 348.1667 | 363.1325 | 332.2 | 322.8333 | 338 | 319.5 | 277.3333 | 320.3333 | 338.6217 | 316.2 | 320.7167 | 354.7241 | 333 | 356.0909 | 327.8333 | 332 | 346.3333 | 333.1667 | 318.5 | 337.2667 | 316.5 | 357.6667 | 315 | 363.8556 | 125.7443 | 330.0513
| Median |  | 356.675 | 339.8 | 334 | 350.5 | 361.3736 | 334.25 | 324 | 339 | 319.5 | 276 | 320 | 338.01 | 315.8 | 319.1 | 353.4591 | 334 | 354.8576 | 326.5 | 331 | 346.5 | 333.5 | 316.5 | 338.6 | 316.5 | 352 | 315 | 365.8 | 127.5745 | 333.32
| Std Dev. |  | 1.4691 | 6.2584 | 12.628 | 5.3448 | 11.6987 | 10.1392 | 5.1543 | 6.7231 | 5.9582 | 6.2183 | 8.0416 | 4.359 | 3.0325 | 14.7917 | 3.7103 | 5.1769 | 11.3083 | 6.1779 | 4.1952 | 8.4538 | 2.7869 | 5.8224 | 10.4542 | 8.4083 | 21.4352 | 5.5498 | 4.4239 | 5.1638 | 7.4323
| Rel.Std.Dev. |  | 0.4117% | 1.853% | 3.721% | 1.535% | 3.222% | 3.052% | 1.597% | 1.989% | 1.865% | 2.242% | 2.51% | 1.287% | 0.959% | 4.612% | 1.046% | 1.555% | 3.176% | 1.884% | 1.264% | 2.441% | 0.8365% | 1.828% | 3.1% | 2.657% | 5.993% | 1.762% | 1.216% | 4.107% | 2.252%
| PDM3 |  | 6.253% | 0.5869% | 1.043% | 3.674% | 8.13% | -1.081% | -3.87% | 0.6464% | -4.862% | -17.42% | -4.614% | 0.8315% | -5.845% | -4.5% | 5.626% | -0.8424% | 6.033% | -2.381% | -1.14% | 3.128% | -0.7928% | -5.16% | 0.428% | -5.756% | 6.503% | -6.202% | 8.345% | -62.56% | -1.72%
| Z-Score (Absolute) |  | 1.21 | 0.24 | 0.32 | 0.77 | 1.52 | 0.04 | 0.51 | 0.25 | 0.68 | 2.82 | 0.64 | 0.28 | 0.85 | 0.62 | 1.1 | 0 | 1.17 | 0.26 | 0.05 | 0.67 | 0.01 | 0.73 | 0.22 | 0.83 | 1.25 | 0.91 | 1.56 | 10.48 | 0.15
| Analytical results for Dy in OREAS 45f (Certified Value 1.49 ppm) | [31 empty] | Certified
Dy | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 18; 10: 23; 13: 24; 14: 33; 15: 38; 17: 39; 18: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*OES; 18: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.4g; 10: 0.5g; 13: 01g; 14: 0.4g; 15: 01g; 17: 0.5g; 18: 0.5g |  | 3
| 1 | 1 | 02: 1.639; 04: 1.5; 05: 1.54; 07: 1.44; 10: 1.36; 13: 1.3808; 14: 1.58; 15: 1.71; 17: 0.48; 18: 1.3591 |  | 1
| 1 | 2 | 02: 1.6002; 04: 1.53; 05: 1.52; 07: 1.43; 10: 1.4; 13: 1.3965; 14: 1.4; 15: 1.6533; 17: 0.395; 18: 1.4267 |  | #N/A
| 1 | 3 | 02: 1.6266; 04: 1.48; 05: 1.54; 07: 1.44; 10: 1.34; 13: 1.4831; 14: 1.28; 15: 1.7067; 17: 0.422; 18: 1.3713 |  | 16
| 1 | 4 | 02: 1.6369; 04: 1.55; 05: 1.51; 07: 1.5; 10: 1.33; 13: 1.4786; 14: 1.63; 15: 1.6933; 17: 0.455; 18: 1.3882 |  | 1.4915
| 1 | 5 | 02: 1.6106; 04: 1.47; 05: 1.5; 07: 1.42; 10: 1.35; 13: 1.4854; 14: 1.41; 15: 1.6933; 17: 0.408; 18: 1.3961 |  | 149
| 1 | 6 | 02: 1.6019; 04: 1.54; 05: 1.45; 07: 1.33; 10: 1.34; 13: 1.4057; 14: 1.6; 15: 1.64; 17: 0.422; 18: 1.3856
| Mean |  | 02: 1.6192; 04: 1.5117; 05: 1.51; 07: 1.4267; 10: 1.3533; 13: 1.4383; 14: 1.4833; 15: 1.6828; 17: 0.43033; 18: 1.3878
| Median |  | 02: 1.6186; 04: 1.515; 05: 1.515; 07: 1.435; 10: 1.345; 13: 1.4421; 14: 1.495; 15: 1.6933; 17: 0.422; 18: 1.3869
| Std Dev. |  | 02: 0.017288; 04: 0.033116; 05: 0.033466; 07: 0.055015; 10: 0.025033; 13: 0.048928; 14: 0.1401; 15: 0.029091; 17: 0.0315; 18: 0.023173
| Rel.Std.Dev. |  | 02: 1.068%; 04: 2.191%; 05: 2.216%; 07: 3.856%; 10: 1.85%; 13: 3.402%; 14: 9.445%; 15: 1.729%; 17: 7.32%; 18: 1.67%
| PDM3 |  | 02: 8.565%; 04: 1.355%; 05: 1.243%; 07: -4.344%; 10: -9.261%; 13: -3.561%; 14: -0.545%; 15: 12.83%; 17: -71.15%; 18: -6.948%
| Z-Score (Absolute) |  | 02: 1.72; 04: 0.55; 05: 0.54; 07: 0.37; 10: 1.17; 13: 0.24; 14: 0.24; 15: 2.42; 17: 11.22; 18: 0.8
| Analytical results for Er in OREAS 45f (Certified Value 0.78 ppm) | [31 empty] | Certified
Er | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 18; 10: 23; 13: 24; 14: 33; 15: 38; 17: 39; 18: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*MS; 18: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.4g; 10: 0.5g; 13: 01g; 14: 0.4g; 15: 01g; 17: 0.5g; 18: 0.5g |  | 3
| 1 | 1 | 02: 0.8613; 04: 0.82; 05: 0.81; 07: 0.74; 10: 0.72; 13: 0.70266; 14: 0.87; 15: 0.91667; 17: 0.369; 18: 0.69272 |  | 1
| 1 | 2 | 02: 0.8403; 04: 0.79; 05: 0.77; 07: 0.73; 10: 0.75; 13: 0.71054; 14: 0.8; 15: 0.85667; 17: 0.341; 18: 0.73366 |  | #N/A
| 1 | 3 | 02: 0.8759; 04: 0.78; 05: 0.83; 07: 0.74; 10: 0.71; 13: 0.73961; 14: 0.75; 15: 0.89667; 17: 0.369; 18: 0.70893 |  | 16
| 1 | 4 | 02: 0.8397; 04: 0.79; 05: 0.8; 07: 0.75; 10: 0.7; 13: 0.76803; 14: 0.87; 15: 0.89; 17: 0.365; 18: 0.70846 |  | 0.78286
| 1 | 5 | 02: 0.841; 04: 0.74; 05: 0.8; 07: 0.74; 10: 0.72; 13: 0.77637; 14: 0.74; 15: 0.88333; 17: 0.377; 18: 0.7364 |  | 150
| 1 | 6 | 02: 0.8394; 04: 0.83; 05: 0.76; 07: 0.67; 10: 0.72; 13: 0.70958; 14: 0.87; 15: 0.88; 17: 0.357; 18: 0.71795
| Mean |  | 02: 0.8496; 04: 0.79167; 05: 0.795; 07: 0.72833; 10: 0.72; 13: 0.73446; 14: 0.81667; 15: 0.88722; 17: 0.363; 18: 0.71635
| Median |  | 02: 0.84065; 04: 0.79; 05: 0.8; 07: 0.74; 10: 0.72; 13: 0.72507; 14: 0.835; 15: 0.88667; 17: 0.367; 18: 0.71344
| Std Dev. |  | 02: 0.015434; 04: 0.031885; 05: 0.025884; 07: 0.029269; 10: 0.016733; 13: 0.031976; 14: 0.06186; 15: 0.019823; 17: 0.012586; 18: 0.016613
| Rel.Std.Dev. |  | 02: 1.817%; 04: 4.028%; 05: 3.256%; 07: 4.019%; 10: 2.324%; 13: 4.354%; 14: 7.575%; 15: 2.234%; 17: 3.467%; 18: 2.319%
| PDM3 |  | 02: 8.526%; 04: 1.125%; 05: 1.551%; 07: -6.965%; 10: -8.029%; 13: -6.182%; 14: 4.319%; 15: 13.33%; 17: -53.63%; 18: -8.495%
| Z-Score (Absolute) |  | 02: 1.3; 04: 0.43; 05: 0.48; 07: 0.52; 10: 0.65; 13: 0.43; 14: 0.81; 15: 1.87; 17: 6.01; 18: 0.7
| Analytical results for Eu in OREAS 45f (Certified Value 0.49 ppm) | [31 empty] | Certified
Eu | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 18; 10: 23; 13: 24; 14: 33; 15: 38; 17: 39; 18: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*MS; 18: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.4g; 10: 0.5g; 13: 01g; 14: 0.4g; 15: 01g; 17: 0.5g; 18: 0.5g |  | 2
| 1 | 1 | 02: 0.5788; 04: 0.49; 05: 0.56; 07: 0.39; 10: 0.42; 13: 0.48358; 14: 0.58; 15: 0.58; 17: 0.241; 18: 0.45871 |  | 1
| 1 | 2 | 02: 0.5701; 04: 0.51; 05: 0.51; 07: 0.38; 10: 0.46; 13: 0.48544; 14: 0.46; 15: 0.55; 17: 0.218; 18: 0.47234 |  | #N/A
| 1 | 3 | 02: 0.5678; 04: 0.48; 05: 0.53; 07: 0.38; 10: 0.42; 13: 0.51768; 14: 0.45; 15: 0.56; 17: 0.241; 18: 0.45528 |  | 16
| 1 | 4 | 02: 0.5806; 04: 0.51; 05: 0.54; 07: 0.39; 10: 0.44; 13: 0.50179; 14: 0.53; 15: 0.56; 17: 0.227; 18: 0.46615 |  | 0.49195
| 1 | 5 | 02: 0.567; 04: 0.47; 05: 0.51; 07: 0.37; 10: 0.43; 13: 0.51532; 14: 0.47; 15: 0.56; 17: 0.239; 18: 0.47334 |  | 151
| 1 | 6 | 02: 0.5626; 04: 0.5; 05: 0.47; 07: 0.36; 10: 0.44; 13: 0.48532; 14: 0.56; 15: 0.55333; 17: 0.233; 18: 0.47323
| Mean |  | 02: 0.57115; 04: 0.49333; 05: 0.52; 07: 0.37833; 10: 0.435; 13: 0.49819; 14: 0.50833; 15: 0.56056; 17: 0.23317; 18: 0.46651
| Median |  | 02: 0.56895; 04: 0.495; 05: 0.52; 07: 0.38; 10: 0.435; 13: 0.49362; 14: 0.5; 15: 0.56; 17: 0.236; 18: 0.46925
| Std Dev. |  | 02: 0.0070778; 04: 0.01633; 05: 0.030984; 07: 0.01169; 10: 0.015166; 13: 0.015671; 14: 0.055648; 15: 0.010417; 17: 0.0092177; 18: 0.0079105
| Rel.Std.Dev. |  | 02: 1.239%; 04: 3.31%; 05: 5.958%; 07: 3.09%; 10: 3.486%; 13: 3.146%; 14: 10.95%; 15: 1.858%; 17: 3.953%; 18: 1.696%
| PDM3 |  | 02: 16.1%; 04: 0.282%; 05: 5.703%; 07: -23.09%; 10: -11.58%; 13: 1.269%; 14: 3.331%; 15: 13.95%; 17: -52.6%; 18: -5.171%
| Z-Score (Absolute) |  | 02: 1.13; 04: 0.04; 05: 0.36; 07: 1.76; 10: 0.91; 13: 0.04; 14: 0.19; 15: 0.97; 17: 3.93; 18: 0.44
| Analytical results for Fe in OREAS 45f (Certified Value 13.69 wt.%) | [31 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 3
| 1 | 1 | 13.7606 | 13.16 | 13.5 | 14.5 | 12.6953 | 14.61 | 13.9 | 11.1775 | 13.4 | 13 | 13.5 | 13.77 | 12.79 | 13.592 | 14.5631 | 14 | 12.6 | 13.8 | 14.15 | 13.2 | 13.85 | 14.5 | 14.33 | 12.9 | 13.5 | 13 | 14.3967 | 9.6514 | 15.2793 |  | 1
| 1 | 2 | 13.4158 | 13.15 | 13.6 | 14.4 | 11.5768 | 14.09 | 13.9 | 10.6866 | 13.39 | 12.9 | 13.2 | 13.91 | 13.07 | 13.968 | 14.4371 | 13.8 | 12.8 | 14 | 14.1 | 13.75 | 14.05 | 14.4 | 13.83 | 12.55 | 13.8 | 12.5 | 14.2967 | 9.5338 | 15.2065 |  | #N/A
| 1 | 3 | 13.3783 | 13.25 | 13.9 | 14.4 | 11.3917 | 14.52 | 14.1 | 10.9579 | 13.35 | 12.7 | 13.9 | 13.74 | 13.01 | 13.503 | 14.3514 | 14.2 | 12.5 | 14 | 13.85 | 14 | 13.95 | 14.3 | 13.53 | 13.25 | 14.5 | 13 | 14.3633 | 9.6773 | 15.6422 |  | 16
| 1 | 4 | 13.395 | 13.26 | 13.9 | 14.3 | 12.4833 | 13.97 | 13.9 | 10.8959 | 13.57 | 13.3 | 13.6 | 13.67 | 13.07 | 13.045 | 14.6272 | 14.2 | 13.3 | 13.95 | 13.95 | 14.4 | 13.95 | 14.1 | 13.57 | 13.3 | 14.2 | 12.5 | 14.3267 | 9.5016 | 15.3533 |  | 13.6892
| 1 | 5 | 13.381 | 13.31 | 13.7 | 14.5 | 12.9654 | 14.01 | 13.75 | 10.9247 | 13.47 | 12.6 | 13.3 | 13.72 | 12.96 | 13.17 | 14.5805 | 14 | 13.3 | 13.85 | 13.9 | 13.4 | 14 | 14.3 | 14.23 | 13.1 | 13.3 | 12.9 | 14.2667 | 9.744 | 15.6839 |  | 152
| 1 | 6 | 13.5696 | 13.24 | 13.6 | 14.6 | 11.916 | 13.63 | 13.65 | 11.091 | 13.21 | 12.7 | 13 | 13.59 | 13.2 | 13.676 | 14.329 | 13.9 | 12.7 | 13.6 | 14.05 | 15.05 | 13.8 | 14.8 | 14.05 | 13.25 | 13.9 | 12.6 | 14.21 | 9.6676 | 15.2364
| Mean |  | 13.4834 | 13.2283 | 13.7 | 14.45 | 12.1714 | 14.1383 | 13.8667 | 10.9556 | 13.3983 | 12.8667 | 13.4167 | 13.7333 | 13.0167 | 13.4923 | 14.4814 | 14.0167 | 12.8667 | 13.8667 | 14 | 13.9667 | 13.9333 | 14.4 | 13.9233 | 13.0583 | 13.8667 | 12.75 | 14.31 | 9.6293 | 15.4003
| Median |  | 13.4054 | 13.245 | 13.65 | 14.45 | 12.1997 | 14.05 | 13.9 | 10.9413 | 13.395 | 12.8 | 13.4 | 13.73 | 13.04 | 13.5475 | 14.5001 | 14 | 12.75 | 13.9 | 14 | 13.875 | 13.95 | 14.35 | 13.94 | 13.175 | 13.85 | 12.75 | 14.3117 | 9.6595 | 15.3163
| Std Dev. |  | 0.15374 | 0.061779 | 0.16733 | 0.10488 | 0.63701 | 0.36706 | 0.15384 | 0.16998 | 0.1204 | 0.2582 | 0.31885 | 0.10708 | 0.13706 | 0.33883 | 0.12636 | 0.16021 | 0.35024 | 0.15384 | 0.11832 | 0.68093 | 0.093095 | 0.23664 | 0.33578 | 0.28882 | 0.44121 | 0.2429 | 0.067363 | 0.092567 | 0.20986
| Rel.Std.Dev. |  | 1.14% | 0.467% | 1.221% | 0.7258% | 5.234% | 2.596% | 1.109% | 1.551% | 0.8986% | 2.007% | 2.377% | 0.7797% | 1.053% | 2.511% | 0.8725% | 1.143% | 2.722% | 1.109% | 0.8452% | 4.875% | 0.6681% | 1.643% | 2.412% | 2.212% | 3.182% | 1.905% | 0.4707% | 0.9613% | 1.363%
| PDM3 |  | -1.504% | -3.367% | 0.07877% | 5.558% | -11.09% | 3.281% | 1.296% | -19.97% | -2.125% | -6.009% | -1.991% | 0.3223% | -4.913% | -1.438% | 5.787% | 2.392% | -6.009% | 1.296% | 2.27% | 2.027% | 1.783% | 5.192% | 1.71% | -4.609% | 1.296% | -6.861% | 4.535% | -29.66% | 12.5%
| Z-Score (Absolute) |  | 0.42 | 0.84 | 0.06 | 1.19 | 2.6 | 0.67 | 0.22 | 4.62 | 0.56 | 1.44 | 0.53 | 0 | 1.19 | 0.4 | 1.25 | 0.47 | 1.44 | 0.22 | 0.44 | 0.39 | 0.33 | 1.11 | 0.32 | 1.12 | 0.22 | 1.64 | 0.96 | 6.83 | 2.78
| Analytical results for Ga in OREAS 45f (Certified Value 20.3 ppm) | [31 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 24.418 | 21 | 20.1 | 24.2 | 19.77 | 22.21 | 20.2 | 36 | 22.9 | 18.4 | 18 | 18.12 | 21.9236 | 18.3 | 22.2 | 20.2 | 18.05 | 19 | 21.7 | 20.4 | 20.6 | 17.75 | 19.4 | 25.9767 | 7.387 | 19.735 |  |  |  |  | 1
| 1 | 2 | 24.1684 | 21.6 | 20.2 | 25.8 | 18.0717 | 21.63 | 19.95 | 36 | 22.6 | 19.6 | 19 | 18.47 | 22.1725 | 17.9 | 21.3 | 21 | 18.7 | 18.3 | 21.7 | 18.6 | 20.8 | 18 | 20 | 24.7167 | 6.478 | 18.9033 |  |  |  |  | #N/A
| 1 | 3 | 24.0891 | 20.4 | 20.8 | 24.2 | 18.7982 | 21.58 | 19.7 | 35 | 22.6 | 18.4 | 19 | 18.41 | 23.1767 | 18 | 20.6 | 21.4 | 18.5 | 17.9 | 21.2 | 20 | 18.6 | 19.85 | 19.1 | 24.2367 | 7.713 | 19.905 |  |  |  |  | 16
| 1 | 4 | 24.4502 | 20.5 | 20.9 | 25.4 | 17.7435 | 22.56 | 20.2 | 36 | 22.1 | 18.7 | 20 | 17.21 | 23.405 | 18 | 20.4 | 21.3 | 17.8 | 18.85 | 21.2 | 21.9 | 19.6 | 19.95 | 19.5 | 25.27 | 7.131 | 20.0731 |  |  |  |  | 20.3281
| 1 | 5 | 23.9588 | 20.5 | 20.8 | 24.9 | 19.4596 | 20.89 | 19.15 | 36 | 22.2 | 18.8 | 20 | 17.71 | 23.2637 | 18 | 19.9 | 21.1 | 18.05 | 18.2 | 21.8 | 17.7 | 20.2 | 19.95 | 20 | 24.7 | 7.636 | 20.0267 |  |  |  |  | 153
| 1 | 6 | 23.9821 | 20 | 20.6 | 25.9 | 18.3071 | 16.95 | 18.7 | 35 | 22.2 | 18.8 | 20 | 17.89 | 22.4788 | 17.5 | 21.7 | 20.8 | 17.2 | 18.35 | 21 | 21.1 | 20.2 | 18.95 | 18.9 | 24.63 | 7.534 | 19.5252
| Mean |  | 24.1777 | 20.6667 | 20.5667 | 25.0667 | 18.6917 | 20.97 | 19.65 | 35.6667 | 22.4333 | 18.7833 | 19.3333 | 17.9683 | 22.7367 | 17.95 | 21.0167 | 20.9667 | 18.05 | 18.4333 | 21.4333 | 19.95 | 20 | 19.075 | 19.4833 | 24.9217 | 7.3132 | 19.6947
| Median |  | 24.1287 | 20.5 | 20.7 | 25.15 | 18.5527 | 21.605 | 19.825 | 36 | 22.4 | 18.75 | 19.5 | 18.005 | 22.8277 | 18 | 20.95 | 21.05 | 18.05 | 18.325 | 21.45 | 20.2 | 20.2 | 19.4 | 19.45 | 24.7083 | 7.4605 | 19.82
| Std Dev. |  | 0.21265 | 0.55737 | 0.33862 | 0.75807 | 0.79939 | 2.0514 | 0.60828 | 0.5164 | 0.31411 | 0.44008 | 0.8165 | 0.4728 | 0.62673 | 0.25884 | 0.86583 | 0.43205 | 0.53104 | 0.41433 | 0.33862 | 1.563 | 0.8 | 1.0059 | 0.4535 | 0.61316 | 0.45804 | 0.43689
| Rel.Std.Dev. |  | 0.8795% | 2.697% | 1.646% | 3.024% | 4.277% | 9.782% | 3.096% | 1.448% | 1.4% | 2.343% | 4.223% | 2.631% | 2.756% | 1.442% | 4.12% | 2.061% | 2.942% | 2.248% | 1.58% | 7.835% | 4% | 5.273% | 2.328% | 2.46% | 6.263% | 2.218%
| PDM3 |  | 18.94% | 1.666% | 1.174% | 23.31% | -8.05% | 3.158% | -3.336% | 75.46% | 10.36% | -7.599% | -4.893% | -11.61% | 11.85% | -11.7% | 3.387% | 3.141% | -11.21% | -9.321% | 5.437% | -1.86% | -1.614% | -6.164% | -4.156% | 22.6% | -64.02% | -3.116%
| Z-Score (Absolute) |  | 2.29 | 0.38 | 0.32 | 2.77 | 0.7 | 0.54 | 0.18 | 8.55 | 1.34 | 0.65 | 0.35 | 1.09 | 1.5 | 1.1 | 0.57 | 0.54 | 1.05 | 0.84 | 0.79 | 0.01 | 0.01 | 0.49 | 0.27 | 2.7 | 6.9 | 0.15
| Analytical results for Gd in OREAS 45f (Certified Value 1.7 ppm) | [31 empty] | Certified
Gd | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 18; 10: 23; 13: 24; 14: 33; 15: 38; 17: 39; 18: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*OES; 18: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.4g; 10: 0.5g; 13: 01g; 14: 0.4g; 15: 01g; 17: 0.5g; 18: 0.5g |  | 3
| 1 | 1 | 02: 2.0775; 04: 1.68; 05: 1.78; 07: 1.58; 10: 1.58; 13: 1.6275; 14: 1.95; 15: 2.1367; 17: 1.55; 18: 1.5132 |  | 1
| 1 | 2 | 02: 2.0315; 04: 1.7; 05: 1.77; 07: 1.57; 10: 1.66; 13: 1.6377; 14: 1.71; 15: 2.0567; 17: 1.64; 18: 1.5537 |  | #N/A
| 1 | 3 | 02: 2.049; 04: 1.61; 05: 1.78; 07: 1.59; 10: 1.58; 13: 1.6748; 14: 1.63; 15: 2.06; 17: 1.69; 18: 1.534 |  | 16
| 1 | 4 | 02: 2.0458; 04: 1.76; 05: 1.81; 07: 1.59; 10: 1.55; 13: 1.7514; 14: 1.89; 15: 2.11; 17: 1.59; 18: 1.5231 |  | 1.7022
| 1 | 5 | 02: 2.0231; 04: 1.62; 05: 1.75; 07: 1.52; 10: 1.59; 13: 1.7727; 14: 1.71; 15: 2.0733; 17: 1.7; 18: 1.5711 |  | 154
| 1 | 6 | 02: 1.9884; 04: 1.69; 05: 1.61; 07: 1.46; 10: 1.59; 13: 1.6207; 14: 1.97; 15: 2.04; 17: 1.64; 18: 1.554
| Mean |  | 02: 2.0359; 04: 1.6767; 05: 1.75; 07: 1.5517; 10: 1.5917; 13: 1.6808; 14: 1.81; 15: 2.0794; 17: 1.635; 18: 1.5415
| Median |  | 02: 2.0387; 04: 1.685; 05: 1.775; 07: 1.575; 10: 1.585; 13: 1.6563; 14: 1.8; 15: 2.0667; 17: 1.64; 18: 1.5439
| Std Dev. |  | 02: 0.029798; 04: 0.055377; 05: 0.071274; 07: 0.051929; 10: 0.03656; 13: 0.065979; 14: 0.14422; 15: 0.036601; 17: 0.057533; 18: 0.021777
| Rel.Std.Dev. |  | 02: 1.464%; 04: 3.303%; 05: 4.073%; 07: 3.347%; 10: 2.297%; 13: 3.925%; 14: 7.968%; 15: 1.76%; 17: 3.519%; 18: 1.413%
| PDM3 |  | 02: 19.61%; 04: -1.498%; 05: 2.81%; 07: -8.842%; 10: -6.492%; 13: -1.256%; 14: 6.335%; 15: 22.16%; 17: -3.946%; 18: -9.438%
| Z-Score (Absolute) |  | 02: 2.25; 04: 0.01; 05: 0.45; 07: 0.8; 10: 0.55; 13: 0.01; 14: 0.83; 15: 2.52; 17: 0.28; 18: 0.86
| Analytical results for Ge in OREAS 45f (Certified Value 0.12 ppm) | [31 empty] | Certified
Ge | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab; 25: Lab |  | 1
| No | No | 02: 02; 04: 04; 05: 05; 07: 10; 10: 14; 13: 21; 14: 25; 15: 27; 17: 28; 18: 29; 19: 30; 20: 31; 21: 34; 23: 35; 24: 36; 25: 39 |  | ppm
|  |  | 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 07: AR*OES/MS; 10: AR*MS; 13: AR*OES; 14: AR*MS; 15: AR*OES/MS; 17: AR*MS; 18: AR*MS; 19: AR*MS; 20: AR*MS; 21: AR*MS; 23: AR*MS; 24: AR*OES/MS; 25: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 0.25g; 05: 0.25g; 07: 01g; 10: 0.5g; 13: 0.15g; 14: 0.25g; 15: 0.25g; 17: 0.5g; 18: 0.5g; 19: 0.5g; 20: 0.5g; 21: 0.25g; 23: 0.5g; 24: 0.5g; 25: 0.5g |  | 2
| 1 | 1 | 02: < 0.5; 04: 0.3; 05: 0.1; 07: 0.65357; 10: 0.1; 13: < 10; 14: < 0.1; 15: 0.3; 17: 0.129; 18: 0.104; 19: 0.14; 20: 0.144; 21: 0.1; 23: 0.099; 24: < 0.05; 25: < 0.005 |  | 1
| 1 | 2 | 02: < 0.5; 04: 0.3; 05: 0.1; 07: 0.60315; 10: 0.097; 13: < 10; 14: < 0.1; 15: 0.3; 17: 0.157; 18: 0.113; 19: 0.12; 20: 0.138; 21: < 0.1; 23: 0.093; 24: < 0.05; 25: < 0.005 |  | 11
| 1 | 3 | 02: < 0.5; 04: 0.3; 05: 0.1; 07: 0.65151; 10: 0.097; 13: < 10; 14: < 0.1; 15: 0.3; 17: 0.139; 18: 0.108; 19: 0.11; 20: 0.14; 21: < 0.1; 23: 0.107; 24: < 0.05; 25: < 0.005 |  | 16
| 1 | 4 | 02: < 0.5; 04: 0.4; 05: 0.1; 07: 0.67919; 10: 0.104; 13: < 10; 14: < 0.1; 15: 0.2; 17: 0.132; 18: 0.113; 19: 0.11; 20: 0.141; 21: 0.1; 23: 0.118; 24: < 0.05; 25: < 0.005 |  | 0.11746
| 1 | 5 | 02: < 0.5; 04: 0.3; 05: 0.1; 07: 0.66988; 10: 0.096; 13: < 10; 14: < 0.1; 15: 0.3; 17: 0.122; 18: 0.114; 19: 0.12; 20: 0.144; 21: < 0.1; 23: 0.11; 24: < 0.05; 25: < 0.005 |  | 155
| 1 | 6 | 02: < 0.5; 04: 0.3; 05: 0.1; 07: 0.63293; 10: 0.097; 13: < 10; 14: < 0.1; 15: 0.3; 17: 0.118; 18: 0.106; 19: 0.12; 20: 0.134; 21: < 0.1; 23: 0.101; 24: < 0.05; 25: < 0.005
| Mean |  | 04: 0.31667; 05: 0.1; 07: 0.64837; 10: 0.0985; 15: 0.28333; 17: 0.13283; 18: 0.10967; 19: 0.12; 20: 0.14017; 21: 0.1; 23: 0.10467
| Median |  | 04: 0.3; 05: 0.1; 07: 0.65254; 10: 0.097; 15: 0.3; 17: 0.1305; 18: 0.1105; 19: 0.12; 20: 0.1405; 21: 0.1; 23: 0.104
| Std Dev. |  | 04: 0.040825; 05: 1.5202e-17; 07: 0.027321; 10: 0.0030166; 15: 0.040825; 17: 0.013963; 18: 0.0042269; 19: 0.010954; 20: 0.0038166; 21: 0; 23: 0.0088694
| Rel.Std.Dev. |  | 04: 12.89%; 05: 1.52e-14%; 07: 4.214%; 10: 3.063%; 15: 14.41%; 17: 10.51%; 18: 3.854%; 19: 9.129%; 20: 2.723%; 21: 0%; 23: 8.474%
| PDM3 |  | 04: 169.6%; 05: -14.86%; 07: 452%; 10: -16.14%; 15: 141.2%; 17: 13.09%; 18: -6.631%; 19: 2.166%; 20: 19.34%; 21: -14.86%; 23: -10.89%
| Z-Score (Absolute) |  | 02: 2.02; 04: NA; 05: NA; 07: 7.98; 10: 0.24; 13: 73.07; 14: 0.97; 15: NA; 17: 0.27; 18: 0.08; 19: 0.08; 20: 0.38; 21: NA; 23: 0.15; 24: 1.34; 25: 1.68
Results from laboratories 4, 5, 27 & 34 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Hf in OREAS 45f (Certified Value 0.93 ppm) | [31 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 18 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.4g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 1.608 | 0.85 | 1.03 | 0.97 | 1.8731 | 0.52 | 0.771 | 0.91 | 0.63 | 1.0744 | 0.945 | 0.71 | 0.759 | 0.848 | 0.81 | 0.852 | 1.05 | 1.08 | 0.868 | 1.28 | 1.2367 | 0.517 | 0.82999 | [7 empty] | 1
| 1 | 2 | 1.5528 | 0.96 | 1.06 | 1 | 1.7689 | 0.47 | 0.715 | 0.9 | 0.7 | 1.0684 | 0.986 | 0.68 | 0.784 | 0.851 | 0.79 | 0.842 | 1.09 | 0.99 | 0.893 | 1.32 | 1.1267 | 0.47 | 0.8542 | [7 empty] | #N/A
| 1 | 3 | 1.6019 | 0.92 | 1.01 | 0.95 | 1.6943 | 0.53 | 0.776 | 0.89 | 0.64 | 1.0992 | 1.09 | 0.68 | 0.759 | 0.865 | 0.74 | 0.87 | 0.96 | 1.04 | 0.985 | 1.26 | 0.96667 | 0.52 | 0.85743 | [7 empty] | 16
| 1 | 4 | 1.5787 | 0.86 | 1.01 | 0.99 | 1.4723 | 0.42 | 0.778 | 0.91 | 0.6 | 1.1569 | 1.06 | 0.74 | 0.703 | 0.833 | 0.77 | 0.883 | 1.14 | 1.06 | 0.996 | 1.25 | 1.1567 | 0.499 | 0.8293 | [7 empty] | 0.92755
| 1 | 5 | 1.5568 | 0.93 | 1.03 | 0.96 | 1.8248 | 0.52 | 0.712 | 0.87 | 0.64 | 1.1299 | 1.05 | 0.75 | 0.745 | 0.838 | 0.76 | 0.911 | 1.03 | 1.05 | 0.954 | 1.26 | 1.1433 | 0.516 | 0.81965 | [7 empty] | 156
| 1 | 6 | 1.5894 | 0.98 | 1.01 | 1.04 | 1.978 | 0.66 | 0.735 | 0.81 | 0.64 | 1.1174 | 0.986 | 0.74 | 0.738 | 0.859 | 0.8 | 0.855 | 1.08 | 1.06 | 0.965 | 1.31 | 1.1433 | 0.505 | 0.81221
| Mean |  | 1.5813 | 0.91667 | 1.025 | 0.985 | 1.7686 | 0.52 | 0.74783 | 0.88167 | 0.64167 | 1.1077 | 1.0195 | 0.71667 | 0.748 | 0.849 | 0.77833 | 0.86883 | 1.0583 | 1.0467 | 0.9435 | 1.28 | 1.1289 | 0.5045 | 0.8338
| Median |  | 1.584 | 0.925 | 1.02 | 0.98 | 1.7969 | 0.52 | 0.753 | 0.895 | 0.64 | 1.1083 | 1.018 | 0.725 | 0.752 | 0.8495 | 0.78 | 0.8625 | 1.065 | 1.055 | 0.9595 | 1.27 | 1.1433 | 0.5105 | 0.82964
| Std Dev. |  | 0.0229 | 0.052409 | 0.019748 | 0.032711 | 0.17391 | 0.08025 | 0.030877 | 0.038166 | 0.032506 | 0.033856 | 0.055381 | 0.031411 | 0.0271 | 0.012149 | 0.026394 | 0.025199 | 0.061128 | 0.030768 | 0.051578 | 0.028983 | 0.08846 | 0.018706 | 0.018307
| Rel.Std.Dev. |  | 1.448% | 5.717% | 1.927% | 3.321% | 9.834% | 15.43% | 4.129% | 4.329% | 5.066% | 3.056% | 5.432% | 4.383% | 3.623% | 1.431% | 3.391% | 2.9% | 5.776% | 2.94% | 5.467% | 2.264% | 7.836% | 3.708% | 2.196%
| PDM3 |  | 70.48% | -1.174% | 10.51% | 6.193% | 90.67% | -43.94% | -19.38% | -4.947% | -30.82% | 19.42% | 9.913% | -22.74% | -19.36% | -8.469% | -16.09% | -6.331% | 14.1% | 12.84% | 1.719% | 38% | 21.71% | -45.61% | -10.11%
| Z-Score (Absolute) |  | 3.16 | 0 | 0.52 | 0.33 | 4.05 | 1.89 | 0.8 | 0.17 | 1.31 | 0.91 | 0.49 | 0.95 | 0.8 | 0.32 | 0.66 | 0.23 | 0.67 | 0.62 | 0.13 | 1.73 | 1.01 | 1.96 | 0.39
| Analytical results for Hg in OREAS 45f (Certified Value 0.031 ppm) | [31 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 04 | 05 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [7 empty] | ppm
|  |  | Various [a] | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [7 empty] | 3
|  |  | see footer | 0.25g | 0.5g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | < 1 | 0.03 | 0.031 | < 5 | < 0.1 | < 0.5 | 0.04 | < 0.01 | 0.03 | 0.033741 | 0.019 | 0.1 | 0.032 | 0.029 | 0.03 | 0.031 | < 0.1 | 0.03 | 0.032 | 0.07 | < 0.05 | 0.026 | 0.02841 | [7 empty] | 1
| 1 | 2 | 0.02 | 0.03 | 0.027 | < 5 | < 0.1 | < 0.5 | 0.051 | < 0.01 | 0.03 | 0.033861 | 0.02 | 0.09 | 0.04 | 0.031 | 0.03 | 0.028 | < 0.1 | 0.03 | 0.028 | 0.04 | < 0.05 | 0.031 | 0.03113 | [7 empty] | 19
| 1 | 3 | 0.03 | 0.03 | 0.029 | < 5 | < 0.1 | < 0.5 | 0.033 | < 0.01 | 0.04 | 0.038126 | 0.052 | 0.07 | 0.035 | 0.031 | 0.04 | 0.027 | < 0.1 | 0.04 | 0.022 | 0.05 | < 0.05 | 0.025 | 0.03212 | [7 empty] | 16
| 1 | 4 | 0.03 | 0.03 | 0.028 | < 5 | < 0.1 | < 0.5 | 0.043 | < 0.01 | 0.03 | 0.03786 | 0.039 | 0.08 | 0.038 | 0.027 | 0.03 | 0.029 | < 0.1 | 0.01 | 0.024 | 0.03 | < 0.05 | 0.031 | 0.02905 | [7 empty] | 0.030726
| 1 | 5 | 0.04 | 0.03 | 0.032 | < 5 | < 0.1 | < 0.5 | 0.027 | < 0.01 | 0.03 | 0.034904 | 0.037 | 0.07 | 0.038 | 0.03 | 0.03 | 0.027 | < 0.1 | 0.02 | 0.023 | 0.04 | < 0.05 | 0.029 | 0.03318 | [7 empty] | 157
| 1 | 6 | 0.04 | 0.03 | 0.029 | < 5 | < 0.1 | < 0.5 | 0.031 | < 0.01 | 0.04 | 0.032383 | 0.024 | 0.08 | 0.035 | 0.035 | 0.03 | 0.028 | < 0.1 | 0.03 | 0.029 | 0.05 | < 0.05 | 0.028 | 0.02821
| Mean |  | 0.032 | 0.03 | 0.029333 |  |  |  | 0.0375 |  | 0.033333 | 0.035146 | 0.031833 | 0.081667 | 0.036333 | 0.0305 | 0.031667 | 0.028333 |  | 0.026667 | 0.026333 | 0.046667 |  | 0.028333 | 0.03035
| Median |  | 0.03 | 0.03 | 0.029 |  |  |  | 0.0365 |  | 0.03 | 0.034383 | 0.0305 | 0.08 | 0.0365 | 0.0305 | 0.03 | 0.028 |  | 0.03 | 0.026 | 0.045 |  | 0.0285 | 0.03009
| Std Dev. |  | 0.0083666 | 0 | 0.0018619 |  |  |  | 0.0088487 |  | 0.005164 | 0.0023481 | 0.013045 | 0.01169 | 0.0028752 | 0.0026646 | 0.0040825 | 0.0015055 |  | 0.010328 | 0.0039328 | 0.013663 |  | 0.0025033 | 0.0020873
| Rel.Std.Dev. |  | 26.15% | 0% | 6.347% |  |  |  | 23.6% |  | 15.49% | 6.681% | 40.98% | 14.31% | 7.913% | 8.736% | 12.89% | 5.314% |  | 38.73% | 14.93% | 29.28% |  | 8.835% | 6.877%
| PDM3 |  | 4.145% | -2.364% | -4.533% |  |  |  | 22.05% |  | 8.485% | 14.38% | 3.603% | 165.8% | 18.25% | -0.7363% | 3.061% | -7.788% |  | -13.21% | -14.3% | 51.88% |  | -7.788% | -1.225%
| Z-Score (Absolute) |  | NA | NA | 0.25 | 332.95 | 2.54 | 29.51 | 0.85 | 3.53 | NA | 0.54 | 0.09 | NA | 0.7 | 0.09 | NA | 0.38 | 2.54 | NA | 0.65 | NA | 0.83 | 0.38 | 0.11
Results from laboratories 4, 5, 23, 27, 30, 34, & 36 were removed due to their 0.01 ppm or greater reading resolution.
[a]: Lab 04: Rnd 01 - AR*MS (0.25g), Rnd 01 - AR*OES (0.25g)
| Analytical results for Ho in OREAS 45f (Certified Value 0.28 ppm) | [31 empty] | Certified
Ho | Round | Replicate | 04: Lab; 05: Lab; 14: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab; 25: Lab | [7 empty] | 1
| No | No | 04: 02; 05: 07; 14: 10; 17: 13; 18: 18; 19: 23; 20: 24; 21: 33; 23: 38; 24: 39; 25: 40 | [7 empty] | ppm
|  |  | 04: AR*MS; 05: AR*OES/MS; 14: AR*OES/MS; 17: AR*OES/MS; 18: AR*MS; 19: AR*MS; 20: AR*OES/MS; 21: AR*MS; 23: AR*MS; 24: AR*MS; 25: AR*MS | [7 empty] | 2
|  |  | 04: 0.5g; 05: 01g; 14: 01g; 17: 25g; 18: 0.4g; 19: 0.5g; 20: 01g; 21: 0.4g; 23: 01g; 24: 0.5g; 25: 0.5g | [7 empty] | 3
| 1 | 1 | 04: < 0.5; 05: 0.29; 14: 0.3736; 17: 0.3; 18: 0.27; 19: 0.25; 20: 0.25887; 21: 0.3; 23: 0.31667; 24: 0.135; 25: 0.25873 | [7 empty] | 1
| 1 | 2 | 04: < 0.5; 05: 0.28; 14: 0.32172; 17: 0.3; 18: 0.27; 19: 0.27; 20: 0.25432; 21: 0.28; 23: 0.31; 24: 0.12; 25: 0.26866 | [7 empty] | #N/A
| 1 | 3 | 04: < 0.5; 05: 0.26; 14: 0.3521; 17: 0.3; 18: 0.27; 19: 0.26; 20: 0.25952; 21: 0.26; 23: 0.31667; 24: 0.136; 25: 0.25704 | [7 empty] | 16
| 1 | 4 | 04: < 0.5; 05: 0.28; 14: 0.34352; 17: 0.3; 18: 0.27; 19: 0.26; 20: 0.27783; 21: 0.31; 23: 0.32; 24: 0.131; 25: 0.25608 | [7 empty] | 0.27807
| 1 | 5 | 04: < 0.5; 05: 0.26; 14: 0.3644; 17: 0.3; 18: 0.27; 19: 0.26; 20: 0.27265; 21: 0.26; 23: 0.32333; 24: 0.134; 25: 0.25798 | [7 empty] | 158
| 1 | 6 | 04: < 0.5; 05: 0.28; 14: 0.36525; 17: 0.29; 18: 0.25; 19: 0.26; 20: 0.25477; 21: 0.31; 23: 0.31333; 24: 0.128; 25: 0.26128
| Mean |  | 05: 0.275; 14: 0.35343; 17: 0.29833; 18: 0.26667; 19: 0.26; 20: 0.26299; 21: 0.28667; 23: 0.31667; 24: 0.13067; 25: 0.25996
| Median |  | 05: 0.28; 14: 0.35825; 17: 0.3; 18: 0.27; 19: 0.26; 20: 0.2592; 21: 0.29; 23: 0.31667; 24: 0.1325; 25: 0.25836
| Std Dev. |  | 05: 0.012247; 14: 0.018826; 17: 0.0040825; 18: 0.008165; 19: 0.0063246; 20: 0.0098522; 21: 0.023381; 23: 0.004714; 24: 0.0059889; 25: 0.004614
| Rel.Std.Dev. |  | 05: 4.454%; 14: 5.327%; 17: 1.368%; 18: 3.062%; 19: 2.433%; 20: 3.746%; 21: 8.156%; 23: 1.489%; 24: 4.583%; 25: 1.775%
| PDM3 |  | 05: -1.104%; 14: 27.1%; 17: 7.288%; 18: -4.1%; 19: -6.498%; 20: -5.421%; 21: 3.092%; 23: 13.88%; 24: -53.01%; 25: -6.512%
| Z-Score (Absolute) |  | 04: 0.67; 05: 0.34; 14: 3.51; 17: 1.28; 18: 0; 19: 0.27; 20: 0.15; 21: 0.81; 23: 2.02; 24: 5.5; 25: 0.27
| Analytical results for I in OREAS 45f (Indicative Value 15.8 ppm) | [31 empty] | Indicative
I | Round | Replicate | 04: Lab | [7 empty] | 1
| No | No | 04: 39 | [7 empty] | ppm
|  |  | 04: AR*MS | [7 empty] | 1
|  |  | 04: 0.5g | [7 empty] | 1
| 1 | 1 | 04: 11 | [7 empty] | 1
| 1 | 2 | 04: 13 | [7 empty] | 12
| 1 | 3 | 04: 16 | [7 empty] | 16
| 1 | 4 | 04: 17 | [7 empty] | 15.8333
| 1 | 5 | 04: 19 | [7 empty] | 31
| 1 | 6 | 04: 19
| Mean |  | 04: 15.8333
| Median |  | 04: 16.5
| Std Dev. |  | 04: 3.2506
| Rel.Std.Dev. |  | 04: 20.53%
| PDM3 |  | 04: 2.22e-13%
Z-Score (Absolute) | bar: 04
| Analytical results for In in OREAS 45f (Certified Value 0.087 ppm) | [31 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 04 | 05 | 07 | 13 | 14 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.25g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.08 | 0.09 | 0.086 | 0.09 | 0.087 | < 5 | 0.08 | < 0.5 | < 10 | 0.09 | 0.088173 | 0.079 | 0.07 | 0.087 | 0.084 | 0.089 | 0.077 | 0.08 | 0.1 | 0.082 | 0.098 | 0.093333 | 0.041 | 0.08596 | [6 empty] | 1
| 1 | 2 | 0.09 | 0.08 | 0.095 | 0.105 | 0.088 | < 5 | 0.08 | < 0.5 | < 10 | 0.08 | 0.093034 | 0.081 | 0.07 | 0.087 | 0.083 | 0.083 | 0.085 | 0.08 | 0.1 | 0.085 | 0.097 | 0.093333 | 0.034 | 0.0859 | [6 empty] | 13
| 1 | 3 | 0.09 | 0.09 | 0.098 | 0.111 | 0.089 | < 5 | 0.08 | < 0.5 | < 10 | 0.09 | 0.092429 | 0.086 | 0.07 | 0.095 | 0.08 | 0.081 | 0.087 | 0.08 | 0.09 | 0.088 | 0.092 | 0.09 | 0.038 | 0.08627 | [6 empty] | 16
| 1 | 4 | 0.08 | 0.09 | 0.088 | 0.097 | 0.083 | < 5 | 0.08 | < 0.5 | < 10 | 0.08 | 0.089878 | 0.082 | 0.07 | 0.089 | 0.088 | 0.081 | 0.089 | 0.08 | 0.1 | 0.104 | 0.099 | 0.09 | 0.038 | 0.08756 | [6 empty] | 0.087199
| 1 | 5 | 0.09 | 0.08 | 0.094 | 0.094 | 0.085 | < 5 | 0.07 | < 0.5 | < 10 | 0.08 | 0.088853 | 0.088 | 0.07 | 0.089 | 0.082 | 0.082 | 0.087 | 0.08 | 0.1 | 0.09 | 0.09 | 0.09 | 0.04 | 0.08666 | [6 empty] | 159
| 1 | 6 | 0.09 | 0.09 | 0.093 | 0.084 | 0.088 | < 5 | 0.07 | < 0.5 | < 10 | 0.08 | 0.086612 | 0.078 | 0.07 | 0.086 | 0.085 | 0.082 | 0.092 | 0.08 | 0.1 | 0.083 | 0.092 | 0.096667 | 0.04 | 0.08606
| Mean |  | 0.086667 | 0.086667 | 0.092333 | 0.096833 | 0.086667 |  | 0.076667 |  |  | 0.083333 | 0.08983 | 0.082333 | 0.07 | 0.088833 | 0.083667 | 0.083 | 0.086167 | 0.08 | 0.098333 | 0.088667 | 0.094667 | 0.092222 | 0.0385 | 0.086402
| Median |  | 0.09 | 0.09 | 0.0935 | 0.0955 | 0.0875 |  | 0.08 |  |  | 0.08 | 0.089366 | 0.0815 | 0.07 | 0.088 | 0.0835 | 0.082 | 0.087 | 0.08 | 0.1 | 0.0865 | 0.0945 | 0.091667 | 0.039 | 0.086165
| Std Dev. |  | 0.005164 | 0.005164 | 0.0045019 | 0.0098675 | 0.0022509 |  | 0.005164 |  |  | 0.005164 | 0.002493 | 0.0039328 | 0 | 0.0032506 | 0.0027325 | 0.0030332 | 0.0050761 | 0 | 0.0040825 | 0.0080911 | 0.0037771 | 0.0027217 | 0.00251 | 0.00063076
| Rel.Std.Dev. |  | 5.958% | 5.958% | 4.876% | 10.19% | 2.597% |  | 6.736% |  |  | 6.197% | 2.775% | 4.777% | 0% | 3.659% | 3.266% | 3.654% | 5.891% | 0% | 4.152% | 9.125% | 3.99% | 2.951% | 6.519% | 0.73%
| PDM3 |  | -0.6108% | -0.6108% | 5.888% | 11.05% | -0.6108% |  | -12.08% |  |  | -4.434% | 3.017% | -5.58% | -19.72% | 1.874% | -4.051% | -4.816% | -1.184% | -8.256% | 12.77% | 1.683% | 8.564% | 5.76% | -55.85% | -0.9147%
| Z-Score (Absolute) |  | 0 | 0 | 0.77 | 1.39 | 0 | 329.12 | 1.36 | 22.27 | 670.07 | 0.45 | 0.43 | 0.59 | 2.27 | 0.3 | 0.41 | 0.5 | 0.07 | 0.91 | 1.59 | 0.27 | 1.09 | 0.76 | 6.57 | 0.04
| Analytical results for Ir in OREAS 45f (Indicative Value 7.22 ppb) | [31 empty] | Indicative
Ir | Round | Replicate | 04: Lab; 05: Lab; 14: Lab | [7 empty] | 1
| No | No | 04: 10; 05: 39; 14: 40 | [7 empty] | ppb
|  |  | 04: AR*OES/MS; 05: AR*MS; 14: AR*MS | [7 empty] | 2
|  |  | 04: 01g; 05: 0.5g; 14: 0.5g | [7 empty] | 2
| 1 | 1 | 04: 12.7517; 05: < 2; 14: 0.75 | [7 empty] | 1
| 1 | 2 | 04: 14.1136; 05: < 2; 14: 1.29 | [7 empty] | 26
| 1 | 3 | 04: 12.955; 05: < 2; 14: 0.99 | [7 empty] | 16
| 1 | 4 | 04: 7.6379; 05: < 2; 14: 0.94 | [7 empty] | 7.2204
| 1 | 5 | 04: 13.8201; 05: < 2; 14: 0.81 | [7 empty] | 32
| 1 | 6 | 04: 19.8564; 05: < 2; 14: 0.73
| Mean |  | 04: 13.5225; 14: 0.91833
| Median |  | 04: 13.3876; 14: 0.875
| Std Dev. |  | 04: 3.9013; 14: 0.20942
| Rel.Std.Dev. |  | 04: 28.85%; 14: 22.8%
| PDM3 |  | 04: 87.28%; 14: -87.28%
| Z-Score (Absolute) |  | 04: 103.4; 05: 0; 14: 0.67
| Analytical results for K in OREAS 45f (Certified Value 0.082 wt.%) | [31 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.10124 | 0.08 | 0.08 | 0.106 | 0.079172 | 0.11 | 0.08 | 0.09 | 0.0555 | 0.07 | 0.07 | 0.07 | 0.078 | 0.10954 | 0.0764 | 0.09 | 0.08 | 0.08 | 0.07 | 0.0805 | 0.07 | 0.07 | 0.1 | 0.07 | 0.14667 | 0.041895 | 0.079979 |  |  |  | 1
| 1 | 2 | 0.10095 | 0.08 | 0.09 | 0.11 | 0.079547 | 0.11 | 0.08 | 0.09 | 0.0555 | 0.07 | 0.07 | 0.07 | 0.081 | 0.11025 | 0.0745 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.06 | 0.07 | 0.09 | 0.07 | 0.1412 | 0.0415 | 0.077311 |  |  |  | #N/A
| 1 | 3 | 0.099765 | 0.09 | 0.09 | 0.104 | 0.074882 | 0.11 | 0.08 | 0.09 | 0.0545 | 0.07 | 0.07 | 0.07 | 0.08 | 0.11743 | 0.0848 | 0.09 | 0.08 | 0.08 | 0.07 | 0.0785 | 0.06 | 0.07 | 0.08 | 0.07 | 0.14027 | 0.043036 | 0.082329 |  |  |  | 16
| 1 | 4 | 0.10025 | 0.09 | 0.09 | 0.109 | 0.08078 | 0.12 | 0.08 | 0.09 | 0.0565 | 0.07 | 0.07 | 0.08 | 0.076 | 0.12062 | 0.0821 | 0.09 | 0.08 | 0.08 | 0.07 | 0.078 | 0.06 | 0.07 | 0.08 | 0.07 | 0.14623 | 0.042496 | 0.083716 |  |  |  | 0.082108
| 1 | 5 | 0.099872 | 0.08 | 0.09 | 0.109 | 0.079706 | 0.1 | 0.08 | 0.09 | 0.0545 | 0.07 | 0.07 | 0.07 | 0.077 | 0.11671 | 0.0846 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.07 | 0.14123 | 0.042546 | 0.084693 |  |  |  | 160
| 1 | 6 | 0.099788 | 0.08 | 0.09 | 0.108 | 0.074139 | 0.07 | 0.08 | 0.09 | 0.0505 | 0.07 | 0.07 | 0.08 | 0.079 | 0.11445 | 0.0764 | 0.09 | 0.08 | 0.08 | 0.07 | 0.081 | 0.07 | 0.07 | 0.09 | 0.07 | 0.13817 | 0.04316 | 0.083434
| Mean |  | 0.10031 | 0.083333 | 0.088333 | 0.10767 | 0.078038 | 0.10333 | 0.08 | 0.09 | 0.0545 | 0.07 | 0.07 | 0.073333 | 0.0785 | 0.11483 | 0.0798 | 0.09 | 0.08 | 0.08 | 0.07 | 0.079667 | 0.065 | 0.07 | 0.086667 | 0.07 | 0.14229 | 0.042439 | 0.081911
| Median |  | 0.10006 | 0.08 | 0.09 | 0.1085 | 0.07936 | 0.11 | 0.08 | 0.09 | 0.055 | 0.07 | 0.07 | 0.07 | 0.0785 | 0.11558 | 0.07925 | 0.09 | 0.08 | 0.08 | 0.07 | 0.08 | 0.065 | 0.07 | 0.085 | 0.07 | 0.14122 | 0.042521 | 0.082882
| Std Dev. |  | 0.00063709 | 0.005164 | 0.0040825 | 0.0022509 | 0.002794 | 0.017512 | 0 | 1.5202e-17 | 0.0020976 | 0 | 0 | 0.005164 | 0.0018708 | 0.0043098 | 0.0045725 | 1.5202e-17 | 0 | 0 | 0 | 0.001169 | 0.0054772 | 0 | 0.008165 | 0 | 0.003409 | 0.00064321 | 0.0027707
| Rel.Std.Dev. |  | 0.6351% | 6.197% | 4.622% | 2.091% | 3.58% | 16.95% | 0% | 1.689e-14% | 3.849% | 0% | 0% | 7.042% | 2.383% | 3.753% | 5.73% | 1.689e-14% | 0% | 0% | 0% | 1.467% | 8.427% | 0% | 9.421% | 0% | 2.396% | 1.516% | 3.383%
| PDM3 |  | 22.17% | 1.493% | 7.582% | 31.13% | -4.957% | 25.85% | -2.567% | 9.612% | -33.62% | -14.75% | -14.75% | -10.69% | -4.394% | 39.86% | -2.81% | 9.612% | -2.567% | -2.567% | -14.75% | -2.973% | -20.84% | -14.75% | 5.553% | -14.75% | 73.3% | -48.31% | -0.24%
| Z-Score (Absolute) |  | 1.37 | 0.22 | 0.56 | 1.87 | 0.13 | 1.57 | 0 | 0.67 | 1.72 | 0.67 | 0.67 | 0.45 | 0.1 | 2.35 | 0.01 | 0.67 | 0 | 0 | 0.67 | 0.02 | 1.01 | 0.67 | 0.45 | 0.67 | 4.2 | 2.53 | 0.13
| Analytical results for La in OREAS 45f (Certified Value 10.7 ppm) | [31 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 11.0191 | 10.1 | 10.1 | 12.2 | 13.8282 | 12.3 | 11.05 | 11 | 11 | 9.6 | 11.8 | 10 | 10.27 | 12.3252 | 10.5 | 10.6 | 11.1 | 10.15 | 11.3 | 10.85 | 12.1 | 9.7 | 10.1 | 13.1 | 10.2 | 14.9967 | 6.779 | 10.3465 |  |  | 1
| 1 | 2 | 10.8986 | 10.4 | 10.4 | 12.5 | 12.4127 | 12.2 | 11.45 | 10 | 11 | 9.5 | 12.2 | 10 | 10.79 | 12.5038 | 10.2 | 10.2 | 11.45 | 10.2 | 10.8 | 10.55 | 9.88 | 9.9 | 9.91 | 13.5 | 10.4 | 14.22 | 6.803 | 10.5835 |  |  | #N/A
| 1 | 3 | 10.7837 | 9.8 | 10.7 | 12.4 | 13.542 | 12.2 | 11.15 | 11 | 11 | 9.67 | 11.9 | 10 | 10.38 | 12.8799 | 11 | 9.9 | 11.55 | 10.3 | 10.6 | 10.75 | 10.3 | 8.7 | 10.75 | 12.9 | 10.5 | 14.4067 | 6.777 | 10.6089 |  |  | 16
| 1 | 4 | 11.0021 | 10.3 | 10.8 | 12.6 | 12.7173 | 12.6 | 11.15 | 11 | 12 | 9.74 | 11.9 | 10 | 10.06 | 13.2963 | 10.9 | 9.9 | 11.3 | 10.25 | 11.1 | 10.6 | 10.4 | 9.5 | 11.15 | 13.1 | 10.1 | 14.78 | 6.771 | 10.8892 |  |  | 10.7474
| 1 | 5 | 10.7984 | 10 | 10.7 | 12 | 13.6896 | 11.9 | 11.15 | 11 | 11 | 9.57 | 12.2 | 10 | 10.39 | 13.3038 | 10.7 | 10.5 | 11.3 | 10.25 | 10.9 | 10.55 | 10.8 | 9.6 | 10.55 | 12.9 | 10.3 | 14.56 | 6.807 | 10.7093 |  |  | 161
| 1 | 6 | 10.7668 | 10.3 | 10.6 | 12.7 | 12.8084 | 10 | 11.1 | 11 | 11 | 8.96 | 12 | 10 | 10.24 | 12.3877 | 10.3 | 10.1 | 11.3 | 10.15 | 10.9 | 10.55 | 11.2 | 9.6 | 10.75 | 13 | 10.2 | 14.41 | 6.74 | 10.7224
| Mean |  | 10.8781 | 10.15 | 10.55 | 12.4 | 13.1663 | 11.8667 | 11.175 | 10.8333 | 11.1667 | 9.5067 | 12 | 10 | 10.355 | 12.7828 | 10.6 | 10.2 | 11.3333 | 10.2167 | 10.9333 | 10.6417 | 10.78 | 9.5 | 10.535 | 13.0833 | 10.2833 | 14.5622 | 6.7795 | 10.6433
| Median |  | 10.8485 | 10.2 | 10.65 | 12.45 | 13.1752 | 12.2 | 11.15 | 11 | 11 | 9.585 | 11.95 | 10 | 10.325 | 12.6919 | 10.6 | 10.15 | 11.3 | 10.225 | 10.9 | 10.575 | 10.6 | 9.6 | 10.65 | 13.05 | 10.25 | 14.485 | 6.778 | 10.6591
| Std Dev. |  | 0.11255 | 0.22583 | 0.25884 | 0.26077 | 0.59173 | 0.94163 | 0.14053 | 0.40825 | 0.40825 | 0.28026 | 0.16733 | 0 | 0.24436 | 0.44448 | 0.32249 | 0.29665 | 0.15384 | 0.060553 | 0.24221 | 0.12813 | 0.78791 | 0.41473 | 0.45842 | 0.22286 | 0.1472 | 0.28308 | 0.02428 | 0.18106
| Rel.Std.Dev. |  | 1.035% | 2.225% | 2.453% | 2.103% | 4.494% | 7.935% | 1.258% | 3.768% | 3.656% | 2.948% | 1.394% | 0% | 2.36% | 3.477% | 3.042% | 2.908% | 1.357% | 0.5927% | 2.215% | 1.204% | 7.309% | 4.366% | 4.351% | 1.703% | 1.431% | 1.944% | 0.3581% | 1.701%
| PDM3 |  | 1.216% | -5.559% | -1.837% | 15.38% | 22.51% | 10.41% | 3.978% | 0.7993% | 3.901% | -11.54% | 11.65% | -6.954% | -3.651% | 18.94% | -1.372% | -5.094% | 5.452% | -4.938% | 1.73% | -0.984% | 0.3031% | -11.61% | -1.977% | 21.73% | -4.318% | 35.49% | -36.92% | -0.9688%
| Z-Score (Absolute) |  | 0.22 | 0.75 | 0.22 | 2.26 | 3.29 | 1.55 | 0.62 | 0.16 | 0.61 | 1.61 | 1.73 | 0.95 | 0.48 | 2.77 | 0.15 | 0.69 | 0.83 | 0.66 | 0.3 | 0.09 | 0.09 | 1.62 | 0.24 | 3.18 | 0.57 | 5.16 | 5.27 | 0.09
| Analytical results for Li in OREAS 45f (Indicative Value 9.19 ppm) | [31 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 11.5152 | 11 | 9 | 15.2 | 31.5373 | 14.5 | 8.9 | 6 | 11 | 6 | < 10 | 5 | 9.9 | 16.5815 | 9.43 | 8 | 7 | 9.1 | 8.7 | 8.9 | 6 | 8.8 | 11.3 | 6 | 17.8 | 0.91304 | 9.273 |  |  |  | 1
| 1 | 2 | 11.2867 | 8 | 10 | 16.7 | 29.9802 | 14.6 | 9 | 6 | 11 | 5 | < 10 | 5 | 10.2 | 16.3034 | 8.75 | 8 | 10.9 | 9.1 | 8.2 | 9 | 7 | 9 | 9.7 | 6 | 16.6333 | 0.72546 | 8.2373 |  |  |  | 21
| 1 | 3 | 11.4748 | 10 | 11 | 15.3 | 30.8872 | 13.4 | 8.4 | 6 | 11 | 5 | < 10 | 6 | 9.9 | 16.9504 | 10.5 | 8 | 10.2 | 9.2 | 7.8 | 8.7 | 7 | 11.5 | 9.5 | 6 | 16.4667 | 0.98875 | 9.4279 |  |  |  | 16
| 1 | 4 | 11.898 | 9 | 11 | 16.3 | 29.678 | 15.6 | 8.9 | 6 | 11 | 6 | < 10 | 6 | 9.4 | 17.0975 | 10 | 8 | 9.4 | 8.8 | 8.5 | 8.4 | 8 | 11.5 | 9.2 | 6 | 17.2 | 0.74354 | 9.3935 |  |  |  | 9.1922
| 1 | 5 | 11.55 | 9 | 10 | 15.7 | 31.3051 | 13.3 | 9.5 | 6 | 11 | 5 | < 10 | 6 | 9.5 | 17.3734 | 10.8 | 8 | 10.7 | 9.1 | 7.9 | 8.5 | 6 | 11.6 | 9.1 | 6 | 17.2 | 1.0543 | 10.0918 |  |  |  | 27
| 1 | 6 | 11.8842 | 10 | 11 | 16.1 | 28.1993 | 8.5 | 9 | 6 | 11 | 5 | < 10 | 6 | 9.8 | 16.6277 | 9.38 | 8 | 8 | 8.5 | 8.4 | 8.2 | 6 | 9.8 | 9.9 | 6 | 16.8 | 1.1583 | 9.7291
| Mean |  | 11.6015 | 9.5 | 10.3333 | 15.8833 | 30.2645 | 13.3167 | 8.95 | 6 | 11 | 5.3333 |  | 5.6667 | 9.7833 | 16.8223 | 9.81 | 8 | 9.3667 | 8.9667 | 8.25 | 8.6167 | 6.6667 | 10.3667 | 9.7833 | 6 | 17.0167 | 0.93056 | 9.3588
| Median |  | 11.5326 | 9.5 | 10.5 | 15.9 | 30.4337 | 13.95 | 8.95 | 6 | 11 | 5 |  | 6 | 9.85 | 16.7891 | 9.715 | 8 | 9.8 | 9.1 | 8.3 | 8.6 | 6.5 | 10.65 | 9.6 | 6 | 17 | 0.9509 | 9.4107
| Std Dev. |  | 0.24211 | 1.0488 | 0.8165 | 0.58793 | 1.2474 | 2.5087 | 0.35071 | 0 | 0 | 0.5164 |  | 0.5164 | 0.29269 | 0.39012 | 0.76749 | 0 | 1.568 | 0.26583 | 0.35071 | 0.30605 | 0.8165 | 1.3216 | 0.80104 | 0 | 0.48523 | 0.17198 | 0.62378
| Rel.Std.Dev. |  | 2.087% | 11.04% | 7.902% | 3.702% | 4.122% | 18.84% | 3.919% | 0% | 0% | 9.682% |  | 9.113% | 2.992% | 2.319% | 7.824% | 0% | 16.74% | 2.965% | 4.251% | 3.552% | 12.25% | 12.75% | 8.188% | 0% | 2.851% | 18.48% | 6.665%
| PDM3 |  | 26.21% | 3.349% | 12.41% | 72.79% | 229.2% | 44.87% | -2.635% | -34.73% | 19.67% | -41.98% |  | -38.35% | 6.431% | 83.01% | 6.721% | -12.97% | 1.898% | -2.453% | -10.25% | -6.261% | -27.47% | 12.78% | 6.431% | -34.73% | 85.12% | -89.88% | 1.812%
| Z-Score (Absolute) |  | 0.94 | 0 | 0.37 | 2.87 | 9.33 | 1.72 | 0.25 | 1.57 | 0.67 | 1.87 | 2.02 | 1.72 | 0.13 | 3.29 | 0.14 | 0.67 | 0.06 | 0.24 | 0.56 | 0.4 | NA | 0.39 | 0.13 | NA | 3.38 | 3.85 | 0.06
Results from laboratories 34 & 37 were removed due to their 1 ppm or greater reading resolution.
| Analytical results for Lu in OREAS 45f (Certified Value 0.097 ppm) | [31 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [15 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 13 | 18 | 23 | 24 | 25 | 27 | 33 | 34 | 38 | 39 | 40 | [15 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [15 empty] | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 25g | 0.4g | 0.5g | 01g | 0.25g | 0.25g | 0.4g | 0.25g | 01g | 0.5g | 0.5g | [15 empty] | 3
| 1 | 1 | < 0.5 | 0.1 | 0.1 | 0.097 | 0.1 | 0.1 | 0.093 | 0.092096 | 0.0989 | 0.07 | 0.11 | 0.09 | 0.19333 | 0.04 | 0.0876 | [15 empty] | 1
| 1 | 2 | < 0.5 | 0.1 | 0.1 | 0.107 | 0.09 | 0.1 | 0.096 | 0.087188 | 0.0946 | 0.07 | 0.1 | 0.09 | 0.24333 | 0.032 | 0.09308 | [15 empty] | #N/A
| 1 | 3 | < 0.5 | 0.09 | 0.1 | 0.097 | 0.1 | 0.1 | 0.092 | 0.092303 | 0.117 | 0.07 | 0.1 | 0.09 | 0.15667 | 0.039 | 0.08942 | [15 empty] | 16
| 1 | 4 | < 0.5 | 0.09 | 0.1 | 0.105 | 0.1 | 0.1 | 0.087 | 0.090071 | 0.104 | 0.07 | 0.11 | 0.1 | 0.23 | 0.039 | 0.08847 | [15 empty] | 0.097035
| 1 | 5 | < 0.5 | 0.1 | 0.1 | 0.093 | 0.1 | 0.1 | 0.091 | 0.092654 | 0.115 | 0.07 | 0.1 | 0.1 | 0.17333 | 0.04 | 0.09087 | [15 empty] | 162
| 1 | 6 | < 0.5 | 0.1 | 0.1 | 0.105 | 0.09 | 0.09 | 0.094 | 0.085726 | 0.0979 | 0.07 | 0.11 | 0.09 | 0.23 | 0.038 | 0.09044
| Mean |  |  | 0.096667 | 0.1 | 0.10067 | 0.096667 | 0.098333 | 0.092167 | 0.090006 | 0.10457 | 0.07 | 0.105 | 0.093333 | 0.20444 | 0.038 | 0.08998
| Median |  |  | 0.1 | 0.1 | 0.101 | 0.1 | 0.1 | 0.0925 | 0.091083 | 0.10145 | 0.07 | 0.105 | 0.09 | 0.21167 | 0.039 | 0.08993
| Std Dev. |  |  | 0.005164 | 1.5202e-17 | 0.0057155 | 0.005164 | 0.0040825 | 0.0030605 | 0.0029297 | 0.0093771 | 0 | 0.0054772 | 0.005164 | 0.035193 | 0.0030332 | 0.0019428
| Rel.Std.Dev. |  |  | 5.342% | 1.52e-14% | 5.678% | 5.342% | 4.152% | 3.321% | 3.255% | 8.968% | 0% | 5.216% | 5.533% | 17.21% | 7.982% | 2.159%
| PDM3 |  |  | -0.3797% | 3.055% | 3.743% | -0.3797% | 1.338% | -5.017% | -7.244% | 7.762% | -27.86% | 8.208% | -3.815% | 110.7% | -60.84% | -7.271%
| Z-Score (Absolute) |  | 15.52 | 0 | 0.34 | 0.4 | 0 | 0.17 | 0.46 | 0.67 | 0.8 | 2.7 | 0.84 | 0.34 | 10.91 | 5.94 | 0.68
| Analytical results for Mg in OREAS 45f (Certified Value 0.152 wt.%) | [31 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.17191 | 0.15 | 0.17 | 0.17 | 0.14425 | 0.19 | 0.13 | 0.167 | 0.19 | 0.129 | 0.15 | 0.14 | 0.13 | 0.151 | 0.1676 | 0.148 | 0.15 | 0.15 | 0.15 | 0.14 | 0.13 | 0.139 | 0.14 | 0.12 | 0.18 | 0.13 | 0.19667 | 0.11454 | 0.16523 |  | 1
| 1 | 2 | 0.17949 | 0.15 | 0.17 | 0.18 | 0.13695 | 0.2 | 0.13 | 0.1585 | 0.19 | 0.127 | 0.15 | 0.15 | 0.14 | 0.153 | 0.16727 | 0.142 | 0.15 | 0.14 | 0.15 | 0.14 | 0.13 | 0.139 | 0.13 | 0.12 | 0.15 | 0.13 | 0.19333 | 0.11596 | 0.16621 |  | #N/A
| 1 | 3 | 0.17107 | 0.16 | 0.18 | 0.17 | 0.12764 | 0.19 | 0.14 | 0.1675 | 0.19 | 0.125 | 0.15 | 0.15 | 0.14 | 0.15 | 0.1744 | 0.154 | 0.15 | 0.15 | 0.15 | 0.14 | 0.13 | 0.144 | 0.13 | 0.13 | 0.15 | 0.13 | 0.19 | 0.11608 | 0.1644 |  | 16
| 1 | 4 | 0.17869 | 0.16 | 0.17 | 0.18 | 0.13737 | 0.2 | 0.13 | 0.1647 | 0.19 | 0.132 | 0.15 | 0.15 | 0.14 | 0.143 | 0.17909 | 0.154 | 0.15 | 0.15 | 0.15 | 0.15 | 0.13 | 0.136 | 0.13 | 0.13 | 0.15 | 0.13 | 0.19667 | 0.11605 | 0.1656 |  | 0.15205
| 1 | 5 | 0.17479 | 0.16 | 0.18 | 0.18 | 0.14597 | 0.19 | 0.13 | 0.1618 | 0.19 | 0.126 | 0.14 | 0.15 | 0.14 | 0.143 | 0.17802 | 0.157 | 0.15 | 0.15 | 0.15 | 0.14 | 0.13 | 0.145 | 0.14 | 0.12 | 0.14 | 0.14 | 0.19667 | 0.11486 | 0.16895 |  | 163
| 1 | 6 | 0.17168 | 0.15 | 0.18 | 0.17 | 0.13576 | 0.16 | 0.13 | 0.1654 | 0.19 | 0.126 | 0.15 | 0.15 | 0.14 | 0.148 | 0.17269 | 0.143 | 0.16 | 0.14 | 0.15 | 0.15 | 0.13 | 0.139 | 0.14 | 0.13 | 0.15 | 0.13 | 0.19333 | 0.11693 | 0.17044
| Mean |  | 0.17461 | 0.155 | 0.175 | 0.175 | 0.13799 | 0.18833 | 0.13167 | 0.16415 | 0.19 | 0.1275 | 0.14833 | 0.14833 | 0.13833 | 0.148 | 0.17318 | 0.14967 | 0.15167 | 0.14667 | 0.15 | 0.14333 | 0.13 | 0.14033 | 0.135 | 0.125 | 0.15333 | 0.13167 | 0.19444 | 0.11574 | 0.16681
| Median |  | 0.17335 | 0.155 | 0.175 | 0.175 | 0.13716 | 0.19 | 0.13 | 0.16505 | 0.19 | 0.1265 | 0.15 | 0.15 | 0.14 | 0.149 | 0.17355 | 0.151 | 0.15 | 0.15 | 0.15 | 0.14 | 0.13 | 0.139 | 0.135 | 0.125 | 0.15 | 0.13 | 0.195 | 0.116 | 0.16591
| Std Dev. |  | 0.0037145 | 0.0054772 | 0.0054772 | 0.0054772 | 0.0065777 | 0.01472 | 0.0040825 | 0.0034251 | 3.0405e-17 | 0.0025884 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0041952 | 0.0050212 | 0.0062823 | 0.0040825 | 0.005164 | 0 | 0.005164 | 0 | 0.0034448 | 0.0054772 | 0.0054772 | 0.013663 | 0.0040825 | 0.0027217 | 0.00088215 | 0.002363
| Rel.Std.Dev. |  | 2.127% | 3.534% | 3.13% | 3.13% | 4.767% | 7.816% | 3.101% | 2.087% | 1.6e-14% | 2.03% | 2.752% | 2.752% | 2.951% | 2.835% | 2.899% | 4.197% | 2.692% | 3.521% | 0% | 3.603% | 0% | 2.455% | 4.057% | 4.382% | 8.91% | 3.101% | 1.4% | 0.7622% | 1.417%
| PDM3 |  | 14.84% | 1.941% | 15.09% | 15.09% | -9.246% | 23.86% | -13.4% | 7.959% | 24.96% | -16.15% | -2.444% | -2.444% | -9.02% | -2.663% | 13.9% | -1.567% | -0.2512% | -3.54% | -1.347% | -5.732% | -14.5% | -7.705% | -11.21% | -17.79% | 0.8449% | -13.4% | 27.88% | -23.88% | 9.706%
| Z-Score (Absolute) |  | 1.12 | 0.28 | 1.14 | 1.14 | 0.44 | 1.71 | 0.71 | 0.67 | 1.78 | 0.89 | 0 | 0 | 0.43 | 0.01 | 1.06 | 0.06 | 0.14 | 0.07 | 0.07 | 0.21 | 0.78 | 0.34 | 0.57 | 0.99 | 0.21 | 0.71 | 1.97 | 1.39 | 0.79
| Analytical results for Mn in OREAS 45f (Certified Value 0.015 wt.%) | [31 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 3
| 1 | 1 | 0.017702 | 0.0185 | 0.0151 | 0.0164 | 0.012792 | 0.0175 | 0.01385 | 0.0143 | 0.028 | 0.0119 | 0.014 | 0.0162 | 0.014 | 0.0143 | 0.020097 | 0.014 | 0.0119 | 0.01455 | 0.014 | 0.0149 | 0.01395 | 0.0145 | 0.0135 | 0.0127 | 0.0175 | 0.0132 | 0.016267 | 0.0069772 | 0.015458 |  | 1
| 1 | 2 | 0.017444 | 0.0187 | 0.0158 | 0.0167 | 0.011195 | 0.0168 | 0.014 | 0.0135 | 0.028 | 0.0118 | 0.013 | 0.0159 | 0.0148 | 0.0147 | 0.020115 | 0.0138 | 0.0123 | 0.01505 | 0.01405 | 0.0158 | 0.0138 | 0.0144 | 0.0152 | 0.0125 | 0.0143 | 0.0131 | 0.015467 | 0.0068576 | 0.015311 |  | #N/A
| 1 | 3 | 0.017756 | 0.0179 | 0.0167 | 0.0162 | 0.01199 | 0.017 | 0.0141 | 0.0141 | 0.028 | 0.0117 | 0.013 | 0.0155 | 0.0142 | 0.0145 | 0.020882 | 0.0151 | 0.0116 | 0.01525 | 0.0147 | 0.0155 | 0.0141 | 0.0141 | 0.0156 | 0.0134 | 0.0153 | 0.0134 | 0.015367 | 0.0073683 | 0.01577 |  | 16
| 1 | 4 | 0.017487 | 0.0179 | 0.0164 | 0.0166 | 0.012415 | 0.0171 | 0.01385 | 0.014 | 0.028 | 0.0122 | 0.013 | 0.0165 | 0.0143 | 0.0138 | 0.021265 | 0.0148 | 0.0125 | 0.01465 | 0.0151 | 0.0142 | 0.014 | 0.0142 | 0.0157 | 0.01385 | 0.0149 | 0.0129 | 0.016133 | 0.0071395 | 0.015828 |  | 0.014642
| 1 | 5 | 0.017542 | 0.0181 | 0.0161 | 0.0168 | 0.013432 | 0.0164 | 0.0135 | 0.0138 | 0.028 | 0.0117 | 0.013 | 0.0166 | 0.0143 | 0.0138 | 0.021639 | 0.0148 | 0.0118 | 0.01475 | 0.01425 | 0.015 | 0.014 | 0.0144 | 0.0136 | 0.0128 | 0.0144 | 0.0131 | 0.0157 | 0.0072741 | 0.016192 |  | 164
| 1 | 6 | 0.017647 | 0.0175 | 0.0161 | 0.0166 | 0.012393 | 0.014 | 0.0135 | 0.0142 | 0.028 | 0.0117 | 0.013 | 0.0154 | 0.0149 | 0.0143 | 0.020464 | 0.014 | 0.0126 | 0.01425 | 0.01445 | 0.0146 | 0.01355 | 0.0146 | 0.0134 | 0.01295 | 0.0147 | 0.0129 | 0.015667 | 0.0073746 | 0.015707
| Mean |  | 0.017596 | 0.0181 | 0.016033 | 0.01655 | 0.012369 | 0.016467 | 0.0138 | 0.013983 | 0.028 | 0.011833 | 0.013167 | 0.016017 | 0.014417 | 0.014233 | 0.020744 | 0.014417 | 0.012117 | 0.01475 | 0.014425 | 0.015 | 0.0139 | 0.014367 | 0.0145 | 0.013033 | 0.015183 | 0.0131 | 0.015767 | 0.0071652 | 0.015711
| Median |  | 0.017594 | 0.018 | 0.0161 | 0.0166 | 0.012404 | 0.0169 | 0.01385 | 0.01405 | 0.028 | 0.01175 | 0.013 | 0.01605 | 0.0143 | 0.0143 | 0.020673 | 0.0144 | 0.0121 | 0.0147 | 0.01435 | 0.01495 | 0.013975 | 0.0144 | 0.0144 | 0.012875 | 0.0148 | 0.0131 | 0.015683 | 0.0072068 | 0.015738
| Std Dev. |  | 0.0001243 | 0.00043818 | 0.00055015 | 0.00021679 | 0.00075275 | 0.0012612 | 0.000251 | 0.00029269 | 0 | 0.00019664 | 0.00040825 | 0.00050365 | 0.00035449 | 0.00036697 | 0.00063012 | 0.00054559 | 0.00040702 | 0.00035777 | 0.00042042 | 0.0005831 | 0.00019748 | 0.00018619 | 0.00111 | 0.00050166 | 0.0011907 | 0.00018974 | 0.00036025 | 0.00021334 | 0.00030732
| Rel.Std.Dev. |  | 0.7064% | 2.421% | 3.431% | 1.31% | 6.086% | 7.659% | 1.819% | 2.093% | 0% | 1.662% | 3.101% | 3.145% | 2.459% | 2.578% | 3.038% | 3.784% | 3.359% | 2.426% | 2.914% | 3.887% | 1.421% | 1.296% | 7.655% | 3.849% | 7.842% | 1.448% | 2.285% | 2.978% | 1.956%
| PDM3 |  | 20.18% | 23.62% | 9.501% | 13.03% | -15.52% | 12.46% | -5.751% | -4.499% | 91.23% | -19.18% | -10.08% | 9.388% | -1.54% | -2.792% | 41.67% | -1.54% | -17.25% | 0.7368% | -1.483% | 2.444% | -5.068% | -1.881% | -0.9707% | -10.99% | 3.696% | -10.53% | 7.68% | -51.06% | 7.3%
| Z-Score (Absolute) |  | 1.61 | 1.87 | 0.82 | 1.08 | 1.05 | 1.04 | 0.32 | 0.22 | 6.91 | 1.32 | 0.64 | 0.81 | 0 | 0.1 | 3.22 | 0 | 1.17 | 0.17 | 0 | 0.29 | 0.27 | 0.03 | 0.04 | 0.71 | 0.39 | 0.67 | 0.68 | 3.69 | 0.65
| Analytical results for Mo in OREAS 45f (Certified Value 1.19 ppm) | [31 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 2
| 1 | 1 | 1.5436 | 1.33 | 1.21 | 1.4 | 2.7808 | 1.17 | 0.94 | 1.16 | < 2 | 1.1 | 1 | 1.05 | < 2 | 1.24 | 1.4542 | 1.17 | 2.67 | 1.08 | 0.86 | 1.06 | 1.29 | 1.4 | 1.47 | 1.15 | 1.7 | < 1 | 1.2667 | 0.718 | 1.2336 |  | 1
| 1 | 2 | 1.574 | 1.22 | 1.19 | 1.4 | 2.8625 | 1.11 | 0.93 | 1.08 | < 2 | 1.1 | 1 | 0.97 | < 2 | 1.22 | 1.4416 | 1.16 | 2.42 | 0.86 | 0.9 | 0.77 | 1.19 | 1.4 | 1.4 | 1.07 | 1.66 | < 1 | 1.2 | 0.79 | 1.2595 |  | 8
| 1 | 3 | 1.545 | 1.33 | 1.18 | 1.4 | 2.8533 | 1.09 | 0.98 | 1.09 | < 2 | 1.1 | 1 | 0.96 | < 2 | 1.19 | 1.5346 | 1.29 | 2.31 | 1.17 | 0.91 | 0.74 | 1.17 | 1.4 | 1.39 | 1.23 | 1.61 | < 1 | 1.1 | 0.91 | 1.3588 |  | 16
| 1 | 4 | 1.5649 | 1.21 | 1.23 | 1.4 | 2.9372 | 1.16 | 0.96 | 1.05 | < 2 | 1.2 | 1 | 0.99 | < 2 | 1.14 | 1.5197 | 1.25 | 2.27 | 1.11 | 0.91 | 0.79 | 1.22 | 1.4 | 1.39 | 1.24 | 1.62 | < 1 | 1.2667 | 0.764 | 1.2865 |  | 1.1928
| 1 | 5 | 1.4724 | 1.28 | 1.17 | 1.4 | 2.8861 | 1.04 | 0.94 | 1.12 | < 2 | 1.1 | 1 | 1 | < 2 | 1.46 | 1.6752 | 1.21 | 2.56 | 1.08 | 0.91 | 0.81 | 1.22 | 1.5 | 1.4 | 1.2 | 1.7 | < 1 | 1.2667 | 0.83 | 1.2466 |  | 165
| 1 | 6 | 1.524 | 1.28 | 1.24 | 1.4 | 2.9114 | 0.99 | 0.91 | 1.13 | < 2 | 1.1 | 1 | 0.98 | < 2 | 1.22 | 1.6169 | 1.23 | 2.53 | 1.07 | 0.89 | 0.79 | 1.21 | 1.4 | 1.41 | 1.22 | 1.61 | < 1 | 1.2333 | 0.908 | 1.2075
| Mean |  | 1.5373 | 1.275 | 1.2033 | 1.4 | 2.8719 | 1.0933 | 0.94333 | 1.105 |  | 1.1167 | 1 | 0.99167 |  | 1.245 | 1.5404 | 1.2183 | 2.46 | 1.0617 | 0.89667 | 0.82667 | 1.2167 | 1.4167 | 1.41 | 1.185 | 1.65 |  | 1.2222 | 0.82 | 1.2654
| Median |  | 1.5443 | 1.28 | 1.2 | 1.4 | 2.8743 | 1.1 | 0.94 | 1.105 |  | 1.1 | 1 | 0.985 |  | 1.22 | 1.5271 | 1.22 | 2.475 | 1.08 | 0.905 | 0.79 | 1.215 | 1.4 | 1.4 | 1.21 | 1.64 |  | 1.25 | 0.81 | 1.2531
| Std Dev. |  | 0.036322 | 0.051672 | 0.028048 | 2.4324e-16 | 0.054336 | 0.069474 | 0.024221 | 0.03937 |  | 0.040825 | 0 | 0.031885 |  | 0.11095 | 0.091309 | 0.04916 | 0.1544 | 0.10534 | 0.019664 | 0.11673 | 0.040825 | 0.040825 | 0.030332 | 0.06473 | 0.042895 |  | 0.065546 | 0.077954 | 0.052755
| Rel.Std.Dev. |  | 2.363% | 4.053% | 2.331% | 1.737e-14% | 1.892% | 6.354% | 2.568% | 3.563% |  | 3.656% | 0% | 3.215% |  | 8.912% | 5.928% | 4.035% | 6.277% | 9.922% | 2.193% | 14.12% | 3.355% | 2.882% | 2.151% | 5.462% | 2.6% |  | 5.363% | 9.507% | 4.169%
| PDM3 |  | 28.89% | 6.894% | 0.8857% | 17.37% | 140.8% | -8.337% | -20.91% | -7.358% |  | -6.38% | -16.16% | -16.86% |  | 4.379% | 29.14% | 2.143% | 106.2% | -10.99% | -24.82% | -30.69% | 2.004% | 18.77% | 18.21% | -0.6513% | 38.33% |  | 2.469% | -31.25% | 6.092%
| Z-Score (Absolute) |  | 1.11 | 0.24 | 0 | 0.65 | 5.53 | 0.36 | 0.86 | 0.33 | 0.67 | 0.29 | 0.67 | 0.7 | 0.67 | 0.14 | 1.12 | 0.05 | 4.17 | 0.47 | 1.02 | 1.25 | 0.04 | 0.71 | 0.69 | 0.06 | 1.48 | 2.33 | 0.06 | 1.27 | 0.21
| Analytical results for Na in OREAS 45f (Certified Value 0.032 wt.%) | [31 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | wt.%
|  |  | AR*OES | Various [a] | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | see footer | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.035581 | 0.03 | 0.04 | 0.03 | 0.034932 | 0.03 | 0.028 | 0.04 | 0.0365 | 0.021 | 0.026 | 0.031 | 0.017752 | 0.0381 | 0.04 | 0.029 | 0.035 | 0.02 | 0.034 | 0.0295 | 0.03 | 0.032 | 0.04 | 0.033333 | 0.027384 | 0.030991 |  |  |  |  | 1
| 1 | 2 | 0.035642 | 0.03 | 0.04 | 0.03 | 0.035487 | 0.03 | 0.029 | 0.04 | 0.036 | 0.021 | 0.026 | 0.031 | 0.017386 | 0.0376 | 0.04 | 0.028 | 0.035 | 0.02 | 0.035 | 0.029 | 0.03 | 0.03 | 0.03 | 0.033333 | 0.027335 | 0.03101 |  |  |  |  | #N/A
| 1 | 3 | 0.035942 | 0.04 | 0.04 | 0.03 | 0.036485 | 0.03 | 0.03 | 0.04 | 0.036 | 0.022 | 0.026 | 0.031 | 0.018938 | 0.0424 | 0.03 | 0.031 | 0.035 | 0.03 | 0.035 | 0.029 | 0.03 | 0.034 | 0.03 | 0.03 | 0.027718 | 0.032093 |  |  |  |  | 16
| 1 | 4 | 0.035866 | 0.03 | 0.04 | 0.03 | 0.034522 | 0.03 | 0.029 | 0.04 | 0.0375 | 0.021 | 0.026 | 0.03 | 0.019754 | 0.0397 | 0.04 | 0.031 | 0.035 | 0.03 | 0.035 | 0.0285 | 0.03 | 0.034 | 0.03 | 0.033333 | 0.027604 | 0.031511 |  |  |  |  | 0.032282
| 1 | 5 | 0.035997 | 0.03 | 0.04 | 0.03 | 0.03508 | 0.03 | 0.029 | 0.04 | 0.036 | 0.021 | 0.026 | 0.03 | 0.018752 | 0.0412 | 0.03 | 0.029 | 0.035 | 0.02 | 0.035 | 0.029 | 0.03 | 0.033 | 0.03 | 0.033333 | 0.02798 | 0.031941 |  |  |  |  | 166
| 1 | 6 | 0.035793 | 0.03 | 0.04 | 0.03 | 0.034152 | 0.03 | 0.028 | 0.04 | 0.0365 | 0.021 | 0.026 | 0.03 | 0.017899 | 0.0382 | 0.04 | 0.028 | 0.033 | 0.03 | 0.034 | 0.03 | 0.03 | 0.034 | 0.03 | 0.03 | 0.027981 | 0.03133
| Mean |  | 0.035804 | 0.031667 | 0.04 | 0.03 | 0.03511 | 0.03 | 0.028833 | 0.04 | 0.036417 | 0.021167 | 0.026 | 0.0305 | 0.018414 | 0.039533 | 0.036667 | 0.029333 | 0.034667 | 0.025 | 0.034667 | 0.029167 | 0.03 | 0.032833 | 0.031667 | 0.032222 | 0.027667 | 0.031479
| Median |  | 0.03583 | 0.03 | 0.04 | 0.03 | 0.035006 | 0.03 | 0.029 | 0.04 | 0.03625 | 0.021 | 0.026 | 0.0305 | 0.018325 | 0.03895 | 0.04 | 0.029 | 0.035 | 0.025 | 0.035 | 0.029 | 0.03 | 0.0335 | 0.03 | 0.033333 | 0.027661 | 0.03142
| Std Dev. |  | 0.00016506 | 0.0040825 | 0 | 0 | 0.00081601 | 0 | 0.00075277 | 0 | 0.00058452 | 0.00040825 | 0 | 0.00054772 | 0.00088827 | 0.0019284 | 0.005164 | 0.0013663 | 0.0008165 | 0.0054772 | 0.0005164 | 0.0005164 | 0 | 0.0016021 | 0.0040825 | 0.0017213 | 0.00028052 | 0.00046302
| Rel.Std.Dev. |  | 0.461% | 12.89% | 0% | 0% | 2.324% | 0% | 2.611% | 0% | 1.605% | 1.929% | 0% | 1.796% | 4.824% | 4.878% | 14.08% | 4.658% | 2.355% | 21.91% | 1.49% | 1.771% | 0% | 4.879% | 12.89% | 5.342% | 1.014% | 1.471%
| PDM3 |  | 10.91% | -1.906% | 23.91% | -7.069% | 8.759% | -7.069% | -10.68% | 23.91% | 12.81% | -34.43% | -19.46% | -5.52% | -42.96% | 22.46% | 13.58% | -9.134% | 7.387% | -22.56% | 7.387% | -9.651% | -7.069% | 1.707% | -1.906% | -0.1856% | -14.3% | -2.487%
| Z-Score (Absolute) |  | 0.91 | NA | NA | NA | 0.77 | NA | 0.56 | NA | 1.04 | 2.18 | 1.16 | 0.21 | 2.76 | 1.7 | NA | 0.45 | 0.67 | NA | 0.67 | 0.49 | NA | 0.29 | NA | 0.16 | 0.81 | 0
Results from laboratories 4, 5, 7, 13, 17, 27, 30, 34 & 36 were removed due to their 0.01 wt.% reading resolution.
[a]: Lab 04: Rnd 01 - AR*MS (0.25g), Rnd 01 - AR*OES (0.25g)
| Analytical results for Nb in OREAS 45f (Indicative Value 0.17 ppm) | [31 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 0.5 | 0.28 | 0.25 | 0.16 | 0.20213 | < 0.05 | 0.111 |  | 0.2 | 0.2 | < 10 | 0.1 | 0.23548 | 0.266 | 0.64 | 0.101 | 0.118 | 0.09 | 0.138 | 0.3 | 0.23 | 0.19 | 0.67 | 0.11333 | 0.095 | 0.12217 |  |  |  |  | 1
| 1 | 2 | < 0.5 | 0.3 | 0.26 | 0.14 | 0.2164 | < 0.05 | 0.076 |  | 0.1 | 0.2 | < 10 | 0.09 | 0.231 | 0.278 | 0.7 | 0.091 | 0.118 | 0.09 | 0.139 | 0.3 | 0.18 | 0.195 | 0.72 | 0.096667 | 0.092 | 0.13103 |  |  |  |  | 22
| 1 | 3 | < 0.5 | 0.31 | 0.21 | 0.16 | 0.23384 | < 0.05 | 0.136 |  | 0.1 | 0.2 | < 10 | 0.1 | 0.19777 | 0.277 | 0.51 | 0.103 | 0.123 | 0.08 | 0.155 | 0.2 | 0.24 | 0.182 | 0.63 | 0.07 | 0.094 | 0.1338 |  |  |  |  | 16
| 1 | 4 | < 0.5 | 0.27 | 0.25 | 0.17 | 0.20302 | < 0.05 | 0.071 |  | 0.1 | 0.2 | < 10 | 0.1 | 0.18639 | 0.252 | 0.64 | 0.098 | 0.121 | 0.08 | 0.167 | 0.3 | 0.26 | 0.198 | 0.64 | 0.11 | 0.089 | 0.12614 |  |  |  |  | 0.16544
| 1 | 5 | < 0.5 | 0.3 | 0.25 | 0.13 | 0.24024 | < 0.05 | 0.133 |  | 0.2 | 0.2 | < 10 | 0.09 | 0.23457 | 0.242 | 0.63 | 0.095 | 0.118 | 0.09 | 0.157 | 0.3 | 0.22 | 0.178 | 0.71 | 0.093333 | 0.092 | 0.11541 |  |  |  |  | 28
| 1 | 6 | < 0.5 | 0.33 | 0.25 | 0.17 | 0.19013 | 0.06 | 0.168 |  | 0.1 | 0.2 | < 10 | 0.1 | 0.21704 | 0.254 | 0.69 | 0.096 | 0.134 | 0.09 | 0.136 | 0.3 | 0.25 | 0.209 | 0.66 | 0.10333 | 0.087 | 0.11102
| Mean |  |  | 0.29833 | 0.245 | 0.155 | 0.21429 | 0.06 | 0.11583 |  | 0.13333 | 0.2 |  | 0.096667 | 0.21704 | 0.2615 | 0.635 | 0.097333 | 0.122 | 0.086667 | 0.14867 | 0.28333 | 0.23 | 0.192 | 0.67167 | 0.097778 | 0.0915 | 0.12326
| Median |  |  | 0.3 | 0.25 | 0.16 | 0.20971 | 0.06 | 0.122 |  | 0.1 | 0.2 |  | 0.1 | 0.22402 | 0.26 | 0.64 | 0.097 | 0.1195 | 0.09 | 0.147 | 0.3 | 0.235 | 0.1925 | 0.665 | 0.1 | 0.092 | 0.12416
| Std Dev. |  |  | 0.02137 | 0.017607 | 0.016432 | 0.019591 |  | 0.037531 |  | 0.05164 | 3.0405e-17 |  | 0.005164 | 0.020752 | 0.014557 | 0.06775 | 0.0043205 | 0.006229 | 0.005164 | 0.012754 | 0.040825 | 0.028284 | 0.011261 | 0.03656 | 0.015587 | 0.0030166 | 0.0088587
| Rel.Std.Dev. |  |  | 7.163% | 7.186% | 10.6% | 9.142% |  | 32.4% |  | 38.73% | 1.52e-14% |  | 5.342% | 9.561% | 5.567% | 10.67% | 4.439% | 5.106% | 5.958% | 8.579% | 14.41% | 12.3% | 5.865% | 5.443% | 15.94% | 3.297% | 7.187%
| PDM3 |  |  | 80.33% | 48.09% | -6.308% | 29.53% | -63.73% | -29.98% |  | -19.4% | 20.89% |  | -41.57% | 31.19% | 58.07% | 283.8% | -41.17% | -26.26% | -47.61% | -10.14% | 71.27% | 39.03% | 16.06% | 306% | -40.9% | -44.69% | -25.49%
| Z-Score (Absolute) |  | 0.56 | 1.02 | 0.51 | 0.36 | 0.21 | 1.55 | 0.73 | NA | 0.57 | 0.08 | 46.32 | 0.92 | 0.24 | 0.67 | 4.27 | 0.91 | 0.67 | 1.01 | 0.42 | 0.88 | 0.37 | 0 | 4.62 | 0.91 | 0.97 | 0.66
| Analytical results for Nd in OREAS 45f (Certified Value 10.1 ppm) | [31 empty] | Certified
Nd | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 10; 07: 13; 10: 18; 13: 23; 14: 24; 15: 33; 17: 38; 18: 39; 19: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*OES/MS; 10: AR*MS; 13: AR*MS; 14: AR*OES/MS; 15: AR*MS; 17: AR*MS; 18: AR*OES; 19: AR*MS |  | 1
|  |  | 02: 0.5g; 04: 01g; 05: 01g; 07: 25g; 10: 0.4g; 13: 0.5g; 14: 01g; 15: 0.4g; 17: 01g; 18: 0.5g; 19: 0.5g |  | 1
| 1 | 1 | 02: 10.954; 04: 10.15; 05: 11.1559; 07: 9.46; 10: 7.45; 13: 8.33; 14: 10.3621; 15: 9.99; 17: 12.1933; 18: 10.237; 19: 8.697 |  | 1
| 1 | 2 | 02: 10.772; 04: 10.51; 05: 10.1244; 07: 9.27; 10: 7.48; 13: 8.69; 14: 10.1712; 15: 9.61; 17: 11.4167; 18: 10.014; 19: 9.0697 |  | #N/A
| 1 | 3 | 02: 10.903; 04: 9.94; 05: 10.8676; 07: 9.36; 10: 7.48; 13: 8.45; 14: 10.6879; 15: 9.56; 17: 11.74; 18: 10.401; 19: 8.9085 |  | 16
| 1 | 4 | 02: 10.91; 04: 10.64; 05: 10.3918; 07: 9.56; 10: 7.66; 13: 8.21; 14: 10.9622; 15: 8.88; 17: 12.0133; 18: 10.214; 19: 8.9676 |  | 10.0694
| 1 | 5 | 02: 10.599; 04: 10.77; 05: 11.285; 07: 9.31; 10: 7.43; 13: 8.51; 14: 10.9628; 15: 8.67; 17: 11.7533; 18: 10.369; 19: 9.0582 |  | 167
| 1 | 6 | 02: 10.689; 04: 10.46; 05: 10.5514; 07: 8.24; 10: 6.89; 13: 8.46; 14: 10.119; 15: 9.48; 17: 11.7033; 18: 10.392; 19: 8.8881
| Mean |  | 02: 10.8045; 04: 10.4117; 05: 10.7293; 07: 9.2; 10: 7.3983; 13: 8.4417; 14: 10.5442; 15: 9.365; 17: 11.8033; 18: 10.2712; 19: 8.9315
| Median |  | 02: 10.8375; 04: 10.485; 05: 10.7095; 07: 9.335; 10: 7.465; 13: 8.455; 14: 10.525; 15: 9.52; 17: 11.7467; 18: 10.303; 19: 8.9381
| Std Dev. |  | 02: 0.14129; 04: 0.31096; 05: 0.45197; 07: 0.48195; 10: 0.26225; 13: 0.16302; 14: 0.38045; 15: 0.49399; 17: 0.26916; 18: 0.1494; 19: 0.13701
| Rel.Std.Dev. |  | 02: 1.308%; 04: 2.987%; 05: 4.212%; 07: 5.239%; 10: 3.545%; 13: 1.931%; 14: 3.608%; 15: 5.275%; 17: 2.28%; 18: 1.455%; 19: 1.534%
| PDM3 |  | 02: 7.3%; 04: 3.399%; 05: 6.553%; 07: -8.634%; 10: -26.53%; 13: -16.17%; 14: 4.715%; 15: -6.996%; 17: 17.22%; 18: 2.003%; 19: -11.3%
| Z-Score (Absolute) |  | 02: 0.4; 04: 0.1; 05: 0.34; 07: 0.8; 10: 2.14; 13: 1.36; 14: 0.2; 15: 0.67; 17: 1.14; 18: 0; 19: 1
| Analytical results for Ni in OREAS 45f (Certified Value 192 ppm) | [31 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 195.17 | 194 | 186 | 227 | 214.6875 | 239.3 | 187.5 | 182 | 191 | 138 | 168 | 190.2 | 156 | 181.9 | 228.7045 | 184 | 225.88 | 196 | 188 | 189 | 185.5 | 221 | 153 | 177 | 160 | 146 | 223 | 51.798 | 188.2852 |  | 1
| 1 | 2 | 192.03 | 197 | 184 | 239 | 212.4856 | 234.3 | 185.5 | 172 | 191 | 134 | 168 | 191.4 | 163 | 189.3 | 230.1585 | 180 | 212.7037 | 197.5 | 184.5 | 194.5 | 186.5 | 181 | 147.1 | 173 | 194 | 149 | 216.9 | 48.937 | 182.3983 |  | #N/A
| 1 | 3 | 194.13 | 194 | 203 | 231 | 211.249 | 230.5 | 190.5 | 181 | 190 | 136 | 162 | 193.7 | 158 | 184.3 | 236.4047 | 198 | 212.9134 | 202 | 191 | 193.5 | 187.5 | 176 | 141.5 | 185.5 | 205 | 151 | 217.6333 | 55.425 | 183.6437 |  | 16
| 1 | 4 | 192.73 | 189 | 203 | 235 | 213.7148 | 252 | 187.5 | 180 | 191 | 137 | 170 | 192.6 | 161 | 170 | 235.2544 | 185 | 218.5236 | 200 | 193 | 195 | 185 | 213 | 153.1 | 187 | 193 | 146 | 223.2 | 52.392 | 182.7806 |  | 191.8644
| 1 | 5 | 181.96 | 197 | 192 | 232 | 215.2994 | 229 | 183.5 | 176 | 188 | 133 | 164 | 193.7 | 159 | 177.8 | 232.2098 | 189 | 217.0952 | 201 | 184.5 | 187 | 185 | 182 | 147.1 | 180.5 | 188 | 145 | 218 | 54.362 | 193.5481 |  | 168
| 1 | 6 | 192.2 | 189 | 185 | 234 | 213.9021 | 186.2 | 181 | 183 | 188 | 126 | 164 | 192.8 | 164 | 181.8 | 231.1549 | 182 | 216.1911 | 194.5 | 180 | 201 | 181 | 203 | 153.1 | 183 | 194 | 145 | 217.4333 | 55.5 | 189.6321
| Mean |  | 191.37 | 193.3333 | 192.1667 | 233 | 213.5564 | 228.55 | 185.9167 | 179 | 189.8333 | 134 | 166 | 192.4 | 160.1667 | 180.85 | 232.3145 | 186.3333 | 217.2178 | 198.5 | 186.8333 | 193.3333 | 185.0833 | 196 | 149.15 | 181 | 189 | 147 | 219.3611 | 53.069 | 186.7147
| Median |  | 192.465 | 194 | 189 | 233 | 213.8085 | 232.4 | 186.5 | 180.5 | 190.5 | 135 | 166 | 192.7 | 160 | 181.85 | 231.6823 | 184.5 | 216.6432 | 198.75 | 186.25 | 194 | 185.25 | 192.5 | 150.05 | 181.75 | 193.5 | 146 | 217.8167 | 53.377 | 185.9644
| Std Dev. |  | 4.7661 | 3.6148 | 8.8412 | 4.0497 | 1.4787 | 22.343 | 3.3529 | 4.1952 | 1.472 | 4.3359 | 3.0984 | 1.3726 | 3.0605 | 6.5167 | 2.9802 | 6.4704 | 4.8314 | 2.9665 | 4.7819 | 4.9363 | 2.223 | 18.8892 | 4.753 | 5.3009 | 15.2578 | 2.4495 | 2.9186 | 2.5408 | 4.4992
| Rel.Std.Dev. |  | 2.491% | 1.87% | 4.601% | 1.738% | 0.6924% | 9.776% | 1.803% | 2.344% | 0.7754% | 3.236% | 1.866% | 0.7134% | 1.911% | 3.603% | 1.283% | 3.473% | 2.224% | 1.494% | 2.559% | 2.553% | 1.201% | 9.637% | 3.187% | 2.929% | 8.073% | 1.666% | 1.33% | 4.788% | 2.41%
| PDM3 |  | -0.2577% | 0.7656% | 0.1576% | 21.44% | 11.31% | 19.12% | -3.1% | -6.705% | -1.059% | -30.16% | -13.48% | 0.2792% | -16.52% | -5.741% | 21.08% | -2.883% | 13.21% | 3.458% | -2.622% | 0.7656% | -3.534% | 2.155% | -22.26% | -5.663% | -1.493% | -23.38% | 14.33% | -72.34% | -2.684%
| Z-Score (Absolute) |  | 0.2 | 0.36 | 0.26 | 3.64 | 2.03 | 3.27 | 0.26 | 0.83 | 0.07 | 4.55 | 1.9 | 0.28 | 2.39 | 0.67 | 3.58 | 0.22 | 2.33 | 0.79 | 0.18 | 0.36 | 0.32 | 0.58 | 3.3 | 0.66 | 0 | 3.47 | 2.51 | 11.25 | 0.19
| Analytical results for P in OREAS 45f (Certified Value 0.022 wt.%) | [31 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 04 | 05 | 07 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.25g | 0.25g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.0232 | 0.03 | 0.023 | 0.0219 | 0.022 | 0.0224 | 0.027 | 0.0075 | 0.024 | 0.02 | 0.0202 | 0.02 | 0.022426 | 0.0232 | 0.023 | 0.023 | 0.022 | 0.022 | 0.021 | 0.0235 | 0.0276 | 0.021 | 0.018 | 0.02 | 0.029867 | 0.0148 | 0.023214 |  |  |  | 1
| 1 | 2 | 0.0226 | 0.03 | 0.023 | 0.022 | 0.022 | 0.021 | 0.027 | 0.007 | 0.024 | 0.021 | 0.0201 | 0.022 | 0.022344 | 0.0221 | 0.023 | 0.023 | 0.023 | 0.023 | 0.022 | 0.0225 | 0.0193 | 0.02 | 0.024 | 0.02 | 0.0293 | 0.0151 | 0.023496 |  |  |  | #N/A
| 1 | 3 | 0.022 | 0.03 | 0.022 | 0.0222 | 0.023 | 0.0218 | 0.026 | 0.007 | 0.024 | 0.021 | 0.0205 | 0.022 | 0.023536 | 0.0235 | 0.021 | 0.023 | 0.024 | 0.023 | 0.021 | 0.022 | 0.02 | 0.022 | 0.029 | 0.02 | 0.029433 | 0.0154 | 0.024049 |  |  |  | 16
| 1 | 4 | 0.0223 | 0.03 | 0.022 | 0.0222 | 0.022 | 0.0218 | 0.027 | 0.0075 | 0.024 | 0.02 | 0.0204 | 0.02 | 0.02384 | 0.0232 | 0.024 | 0.022 | 0.021 | 0.023 | 0.022 | 0.022 | 0.0182 | 0.021 | 0.02 | 0.02 | 0.029833 | 0.0152 | 0.023509 |  |  |  | 0.022145
| 1 | 5 | 0.0223 | 0.03 | 0.023 | 0.0217 | 0.022 | 0.0211 | 0.027 | 0.007 | 0.024 | 0.021 | 0.0202 | 0.02 | 0.023572 | 0.0234 | 0.023 | 0.022 | 0.023 | 0.022 | 0.021 | 0.023 | 0.0282 | 0.021 | 0.028 | 0.02 | 0.029333 | 0.0146 | 0.02384 |  |  |  | 169
| 1 | 6 | 0.0222 | 0.03 | 0.022 | 0.0207 | 0.021 | 0.0223 | 0.027 | 0.0075 | 0.024 | 0.021 | 0.0209 | 0.021 | 0.023421 | 0.0228 | 0.024 | 0.022 | 0.023 | 0.025 | 0.021 | 0.024 | 0.0278 | 0.021 | 0.022 | 0.02 | 0.028967 | 0.0145 | 0.023711
| Mean |  | 0.022433 | 0.03 | 0.0225 | 0.021783 | 0.022 | 0.021733 | 0.026833 | 0.00725 | 0.024 | 0.020667 | 0.020383 | 0.020833 | 0.02319 | 0.023033 | 0.023 | 0.0225 | 0.022667 | 0.023 | 0.021333 | 0.022833 | 0.023517 | 0.021 | 0.0235 | 0.02 | 0.029456 | 0.014933 | 0.023637
| Median |  | 0.0223 | 0.03 | 0.0225 | 0.02195 | 0.022 | 0.0218 | 0.027 | 0.00725 | 0.024 | 0.021 | 0.0203 | 0.0205 | 0.023478 | 0.0232 | 0.023 | 0.0225 | 0.023 | 0.023 | 0.021 | 0.02275 | 0.0238 | 0.021 | 0.023 | 0.02 | 0.029383 | 0.01495 | 0.02361
| Std Dev. |  | 0.00042269 | 0 | 0.00054772 | 0.00056362 | 0.00063246 | 0.00058538 | 0.00040825 | 0.00027386 | 3.8006e-18 | 0.0005164 | 0.00029269 | 0.00098319 | 0.00063864 | 0.0005164 | 0.0010954 | 0.00054772 | 0.0010328 | 0.0010954 | 0.0005164 | 0.0008165 | 0.0048035 | 0.00063246 | 0.0043704 | 0 | 0.00034362 | 0.0003559 | 0.00029356
| Rel.Std.Dev. |  | 1.884% | 0% | 2.434% | 2.587% | 2.875% | 2.693% | 1.521% | 3.777% | 1.584e-14% | 2.499% | 1.436% | 4.719% | 2.754% | 2.242% | 4.763% | 2.434% | 4.556% | 4.763% | 2.421% | 3.576% | 20.43% | 3.012% | 18.6% | 0% | 1.167% | 2.383% | 1.242%
| PDM3 |  | 1.301% | 35.47% | 1.602% | -1.635% | -0.6561% | -1.86% | 21.17% | -67.26% | 8.375% | -6.677% | -7.956% | -5.924% | 4.717% | 4.01% | 3.859% | 1.602% | 2.354% | 3.859% | -3.667% | 3.107% | 6.193% | -5.172% | 6.117% | -9.687% | 33.01% | -32.57% | 6.734%
| Z-Score (Absolute) |  | 0.04 | NA | 0 | 0.48 | 0.34 | 0.52 | 2.92 | 10.28 | 1.01 | 1.24 | 1.43 | 1.12 | 0.47 | 0.36 | 0.34 | 0 | 0.11 | 0.34 | 0.79 | 0.22 | 0.69 | 1.01 | 0.67 | NA | 4.69 | 5.1 | 0.77
Results from laboratories 5 & 37 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 45f (Certified Value 12.4 ppm) | [31 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 0.5g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 2
| 1 | 1 | 13.8906 | 13 | 12.2 | 12.9 | 17.1994 | 14 | 12.7 | 12 | 13 | 13 | 12 | 12.7 | 12 | 11.8 | 12.7812 | 12.3 | 14 | 11.3 | 11.65 | 14.8 | 12.6 | 13 | 9.4 | 11.5 | 12.8 | 17 | 12.7667 | 10.6 | 11.5279 |  | 1
| 1 | 2 | 13.4765 | 11 | 12.4 | 13 | 15.681 | 13.8 | 12.65 | 14 | 10 | 13 | 12 | 12.9 | 13 | 12.3 | 12.4243 | 12.2 | 15.7 | 12.5 | 11.7 | 12.5 | 12.45 | 12 | 9.5 | 11.25 | 12.9 | 14 | 12.8 | 10.2 | 11.4923 |  | 9
| 1 | 3 | 13.8617 | 12 | 12.6 | 12.2 | 17.4282 | 14.2 | 12.65 | 13 | 11 | 13 | 12 | 12.5 | 12 | 11.7 | 13.029 | 12.3 | 14.5 | 12.55 | 11.95 | 12.2 | 12.4 | 13 | 9 | 12.25 | 12.4 | 17 | 12.7667 | 10.4 | 11.7993 |  | 16
| 1 | 4 | 13.8669 | 12 | 12.5 | 12.8 | 15.7445 | 13.8 | 12.6 | 12 | 12 | 13 | 12 | 12.4 | 13 | 11.6 | 12.531 | 12.4 | 14.5 | 12.25 | 11.75 | 13.3 | 12.6 | 12 | 9 | 12.65 | 12.6 | 17 | 12.9667 | 10.2 | 11.4972 |  | 12.3904
| 1 | 5 | 13.3138 | 13 | 12.5 | 12.2 | 18.5545 | 13.8 | 12.3 | 13 | 11 | 13 | 12 | 12.4 | 12 | 11.8 | 13.0141 | 12.4 | 13.2 | 12.45 | 11.6 | 12.2 | 12.45 | 13 | 9.4 | 12.15 | 12.6 | 17 | 12.7333 | 10.7 | 11.6859 |  | 170
| 1 | 6 | 13.5424 | 11 | 12.5 | 12.9 | 16.9503 | 12.8 | 12.4 | 13 | 12 | 12 | 12 | 12.4 | 11 | 11.8 | 12.6567 | 12.1 | 14 | 12.1 | 11.65 | 12.6 | 12.4 | 12 | 9.4 | 12.3 | 12.7 | 16 | 12.6667 | 10.4 | 11.6705
| Mean |  | 13.6587 | 12 | 12.45 | 12.6667 | 16.9263 | 13.7333 | 12.55 | 12.8333 | 11.5 | 12.8333 | 12 | 12.55 | 12.1667 | 11.8333 | 12.7394 | 12.2833 | 14.3167 | 12.1917 | 11.7167 | 12.9333 | 12.4833 | 12.5 | 9.2833 | 12.0167 | 12.6667 | 16.3333 | 12.7833 | 10.4167 | 11.6122
| Median |  | 13.7021 | 12 | 12.5 | 12.85 | 17.0749 | 13.8 | 12.625 | 13 | 11.5 | 13 | 12 | 12.45 | 12 | 11.8 | 12.719 | 12.3 | 14.25 | 12.35 | 11.675 | 12.55 | 12.45 | 12.5 | 9.4 | 12.2 | 12.65 | 17 | 12.7667 | 10.4 | 11.5992
| Std Dev. |  | 0.24658 | 0.89443 | 0.13784 | 0.36697 | 1.0886 | 0.48442 | 0.16125 | 0.75277 | 1.0488 | 0.40825 | 0 | 0.20736 | 0.75277 | 0.24221 | 0.24927 | 0.1169 | 0.82805 | 0.46842 | 0.12517 | 0.99933 | 0.093095 | 0.54772 | 0.22286 | 0.53072 | 0.17512 | 1.2111 | 0.10055 | 0.20412 | 0.12532
| Rel.Std.Dev. |  | 1.805% | 7.454% | 1.107% | 2.897% | 6.432% | 3.527% | 1.285% | 5.866% | 9.12% | 3.181% | 0% | 1.652% | 6.187% | 2.047% | 1.957% | 0.9517% | 5.784% | 3.842% | 1.068% | 7.727% | 0.7458% | 4.382% | 2.401% | 4.417% | 1.383% | 7.415% | 0.7866% | 1.96% | 1.079%
| PDM3 |  | 10.24% | -3.151% | 0.4807% | 2.229% | 36.61% | 10.84% | 1.288% | 3.574% | -7.187% | 3.574% | -3.151% | 1.288% | -1.806% | -4.496% | 2.816% | -0.8645% | 15.55% | -1.604% | -5.438% | 4.382% | 0.7497% | 0.8842% | -25.08% | -3.017% | 2.229% | 31.82% | 3.171% | -15.93% | -6.281%
| Z-Score (Absolute) |  | 1.8 | 0.78 | 0.08 | 0.26 | 6.89 | 1.92 | 0.08 | 0.52 | 1.56 | 0.52 | 0.78 | 0.08 | 0.52 | 1.04 | 0.37 | 0.34 | 2.83 | 0.48 | 1.22 | 0.67 | 0.03 | 0 | 5.01 | 0.75 | 0.26 | 5.97 | 0.44 | 3.24 | 1.38
| Analytical results for Pd in OREAS 45f (Certified Value 39.9 ppb) | [31 empty] | Certified
Pd | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 13; 04: 14; 05: 23; 07: 24; 10: 28; 13: 29; 14: 31; 15: 35; 17: 39; 18: 40 |  | ppb
|  |  | 02: AR*OES/MS; 04: AR*MS; 05: AR*MS; 07: AR*OES/MS; 10: AR*MS; 13: AR*MS; 14: AR*MS; 15: AR*MS; 17: AR*MS; 18: AR*MS |  | 1
|  |  | 02: 25g; 04: 0.5g; 05: 0.5g; 07: 01g; 10: 0.5g; 13: 0.5g; 14: 0.5g; 15: 0.5g; 17: 0.5g; 18: 0.5g |  | 1
| 1 | 1 | 02: 43; 04: 41; 05: 38; 07: 45.7191; 10: 39; 13: 38; 14: 44; 15: 38; 17: 21; 18: 37.42 |  | 1
| 1 | 2 | 02: 40; 04: 42; 05: 39; 07: 42.1922; 10: 41; 13: 38; 14: 40; 15: 37; 17: 18; 18: 37.82 |  | 17
| 1 | 3 | 02: 37; 04: 38; 05: 40; 07: 52.4536; 10: 42; 13: 43; 14: 38; 15: 41; 17: 19; 18: 40.69 |  | 16
| 1 | 4 | 02: 41; 04: 39; 05: 37; 07: 46.0103; 10: 43; 13: 42; 14: 43; 15: 46; 17: 18; 18: 36.9 |  | 39.9213
| 1 | 5 | 02: 41; 04: 42; 05: 39; 07: 50.7357; 10: 43; 13: 41; 14: 37; 15: 41; 17: 19; 18: 39.38 |  | 171
| 1 | 6 | 02: 28; 04: 38; 05: 40; 07: 46.2467; 10: 37; 13: 38; 14: 37; 15: 42; 17: 16; 18: 39.61
| Mean |  | 02: 38.3333; 04: 40; 05: 38.8333; 07: 47.2263; 10: 40.8333; 13: 40; 14: 39.8333; 15: 40.8333; 17: 18.5; 18: 38.6367
| Median |  | 02: 40.5; 04: 40; 05: 39; 07: 46.1285; 10: 41.5; 13: 39.5; 14: 39; 15: 41; 17: 18.5; 18: 38.6
| Std Dev. |  | 02: 5.4283; 04: 1.8974; 05: 1.169; 07: 3.7335; 10: 2.4014; 13: 2.2804; 14: 3.0605; 15: 3.1885; 17: 1.6432; 18: 1.4751
| Rel.Std.Dev. |  | 02: 14.16%; 04: 4.743%; 05: 3.01%; 07: 7.905%; 10: 5.881%; 13: 5.701%; 14: 7.683%; 15: 7.809%; 17: 8.882%; 18: 3.818%
| PDM3 |  | 02: -3.978%; 04: 0.1973%; 05: -2.725%; 07: 18.3%; 10: 2.285%; 13: 0.1973%; 14: -0.2202%; 15: 2.285%; 17: -53.66%; 18: -3.218%
| Z-Score (Absolute) |  | 02: 1.07; 04: 0.06; 05: 0.73; 07: 4.93; 10: 0.62; 13: 0.06; 14: 0.06; 15: 0.62; 17: 14.44; 18: 0.86
| Analytical results for Pr in OREAS 45f (Certified Value 2.63 ppm) | [31 empty] | Certified
Pr | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 18; 10: 23; 13: 24; 14: 33; 15: 38; 17: 39; 18: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*OES; 18: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.4g; 10: 0.5g; 13: 01g; 14: 0.4g; 15: 01g; 17: 0.5g; 18: 0.5g |  | 2
| 1 | 1 | 02: 2.8865; 04: 2.752; 05: 2.77; 07: 1.96; 10: 2.22; 13: 2.6978; 14: 3.06; 15: 3.3533; 17: 0.425; 18: 2.2787 |  | 1
| 1 | 2 | 02: 2.8722; 04: 2.843; 05: 2.75; 07: 1.97; 10: 2.34; 13: 2.7615; 14: 2.46; 15: 3.15; 17: 0.263; 18: 2.3502 |  | #N/A
| 1 | 3 | 02: 2.8532; 04: 2.658; 05: 2.72; 07: 1.98; 10: 2.26; 13: 2.8306; 14: 2.55; 15: 3.28; 17: 0.302; 18: 2.336 |  | 16
| 1 | 4 | 02: 2.9031; 04: 2.832; 05: 2.78; 07: 2.01; 10: 2.2; 13: 2.8721; 14: 2.69; 15: 3.3167; 17: 0.502; 18: 2.3607 |  | 2.6333
| 1 | 5 | 02: 2.8584; 04: 2.597; 05: 2.67; 07: 1.97; 10: 2.26; 13: 2.8566; 14: 2.72; 15: 3.2333; 17: 0.393; 18: 2.3707 |  | 172
| 1 | 6 | 02: 2.8692; 04: 2.822; 05: 2.31; 07: 1.85; 10: 2.27; 13: 2.6674; 14: 2.88; 15: 3.18; 17: 0.211; 18: 2.349
| Mean |  | 02: 2.8738; 04: 2.7507; 05: 2.6667; 07: 1.9567; 10: 2.2583; 13: 2.781; 14: 2.7267; 15: 3.2522; 17: 0.34933; 18: 2.3409
| Median |  | 02: 2.8707; 04: 2.787; 05: 2.735; 07: 1.97; 10: 2.26; 13: 2.796; 14: 2.705; 15: 3.2567; 17: 0.3475; 18: 2.3496
| Std Dev. |  | 02: 0.018466; 04: 0.1024; 05: 0.17918; 07: 0.055015; 10: 0.048339; 13: 0.085641; 14: 0.21833; 15: 0.078956; 17: 0.10933; 18: 0.032633
| Rel.Std.Dev. |  | 02: 0.6426%; 04: 3.723%; 05: 6.719%; 07: 2.812%; 10: 2.14%; 13: 3.079%; 14: 8.007%; 15: 2.428%; 17: 31.3%; 18: 1.394%
| PDM3 |  | 02: 9.132%; 04: 4.458%; 05: 1.268%; 07: -25.69%; 10: -14.24%; 13: 5.61%; 14: 3.546%; 15: 23.5%; 17: -86.73%; 18: -11.1%
| Z-Score (Absolute) |  | 02: 0.45; 04: 0.14; 05: 0.08; 07: 1.87; 10: 1.11; 13: 0.21; 14: 0.08; 15: 1.41; 17: 5.94; 18: 0.9
| Analytical results for Pt in OREAS 45f (Certified Value 36.7 ppb) | [31 empty] | Certified
Pt | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab |  | 1
| No | No | 02: 10; 04: 12; 05: 13; 07: 14; 10: 23; 13: 24; 14: 28; 15: 29; 17: 31; 18: 35; 19: 39; 20: 40 |  | ppb
|  |  | 02: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*MS; 18: AR*MS; 19: AR*MS; 20: AR*MS |  | 1
|  |  | 02: 01g; 04: 40g; 05: 25g; 07: 0.5g; 10: 0.5g; 13: 01g; 14: 0.5g; 15: 0.5g; 17: 0.5g; 18: 0.5g; 19: 0.5g; 20: 0.5g |  | 1
| 1 | 1 | 02: 752.7501; 04: 35; 05: 38; 07: 38; 10: 36; 13: 35.7792; 14: 33; 15: 36; 17: 33; 18: 38; 19: 15; 20: 33.55 |  | 1
| 1 | 2 | 02: 775.3365; 04: 35; 05: 39; 07: 39; 10: 36; 13: 36.1855; 14: 38; 15: 38; 17: 34; 18: 39; 19: 12; 20: 37.53 |  | 10
| 1 | 3 | 02: 880.5163; 04: 35; 05: 40; 07: 36; 10: 35; 13: 36.702; 14: 39; 15: 38; 17: 31; 18: 39; 19: 15; 20: 35.95 |  | 16
| 1 | 4 | 02: 830.4998; 04: 35; 05: 39; 07: 41; 10: 33; 13: 35.7389; 14: 38; 15: 39; 17: 35; 18: 40; 19: 13; 20: 36.08 |  | 36.653
| 1 | 5 | 02: 791.2316; 04: 35; 05: 36; 07: 38; 10: 35; 13: 36.8422; 14: 38; 15: 35; 17: 34; 18: 40; 19: 14; 20: 38.13 |  | 173
| 1 | 6 | 02: 751.6981; 04: 30; 05: 36; 07: 37; 10: 35; 13: 37.1706; 14: 37; 15: 38; 17: 35; 18: 36; 19: 13; 20: 36.52
| Mean |  | 02: 797.0054; 04: 34.1667; 05: 38; 07: 38.1667; 10: 35; 13: 36.403; 14: 37.1667; 15: 37.3333; 17: 33.6667; 18: 38.6667; 19: 13.6667; 20: 36.2933
| Median |  | 02: 783.284; 04: 35; 05: 38.5; 07: 38; 10: 35; 13: 36.4437; 14: 38; 15: 38; 17: 34; 18: 39; 19: 13.5; 20: 36.3
| Std Dev. |  | 02: 50.2197; 04: 2.0412; 05: 1.6733; 07: 1.7224; 10: 1.0954; 13: 0.59143; 14: 2.137; 15: 1.5055; 17: 1.5055; 18: 1.5055; 19: 1.2111; 20: 1.5902
| Rel.Std.Dev. |  | 02: 6.301%; 04: 5.974%; 05: 4.403%; 07: 4.513%; 10: 3.13%; 13: 1.625%; 14: 5.75%; 15: 4.033%; 17: 4.472%; 18: 3.894%; 19: 8.861%; 20: 4.381%
| PDM3 |  | 02: 2074%; 04: -6.783%; 05: 3.675%; 07: 4.13%; 10: -4.51%; 13: -0.6819%; 14: 1.402%; 15: 1.856%; 17: -8.148%; 18: 5.494%; 19: -62.71%; 20: -0.9812%
| Z-Score (Absolute) |  | 02: 323.76; 04: 1.12; 05: 0.52; 07: 0.59; 10: 0.76; 13: 0.16; 14: 0.16; 15: 0.23; 17: 1.33; 18: 0.8; 19: 9.85; 20: 0.21
| Analytical results for Rb in OREAS 45f (Certified Value 14.4 ppm) | [31 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 18 | 19 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.4g | 40g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 20.1236 | 15.7 | 14.7 | 20.1 | 14.7747 | 21.57 | 14.5 | 8.82 | 13.3 | 13.86 | 22.6997 | 12.7 | 13.7 | 15.9 | 13.95 | 14.8 | 14.8 | 15.1 | 12.3 | 14.1 | 15.4 | 29.44 | 4.256 | 14.3079 | [6 empty] | 1
| 1 | 2 | 20.0771 | 16.1 | 15.2 | 21 | 12.7293 | 21.35 | 14.8 | 8.46 | 13 | 14.51 | 23.0377 | 13.1 | 14.2 | 16.1 | 13.85 | 13.2 | 14.6 | 15.5 | 12.2 | 13.8 | 16 | 27.5967 | 3.852 | 13.5326 | [6 empty] | 2
| 1 | 3 | 19.9696 | 15.4 | 16 | 20.1 | 14.0181 | 20.09 | 14.9 | 8.54 | 13.4 | 14.18 | 23.9618 | 15.1 | 13.6 | 16.9 | 14.2 | 12.7 | 14.65 | 14.6 | 10.5 | 15.1 | 15 | 26.6 | 4.421 | 14.9382 | [6 empty] | 16
| 1 | 4 | 20.0476 | 15.3 | 15.9 | 20.8 | 13.2574 | 22.89 | 14.6 | 8.81 | 12.6 | 13.44 | 23.904 | 14.2 | 13.5 | 16.25 | 14.15 | 14.1 | 14.5 | 15.9 | 12.2 | 15.4 | 14.9 | 28.87 | 4.116 | 15.2969 | [6 empty] | 14.3727
| 1 | 5 | 19.9926 | 15.4 | 15.8 | 20.1 | 14.6328 | 19.73 | 14 | 8.19 | 12.5 | 13.56 | 24.0237 | 14.8 | 13.6 | 16.35 | 13.8 | 13.8 | 13.85 | 15.3 | 12.2 | 14.65 | 15.4 | 27.7933 | 4.286 | 15.2557 | [6 empty] | 174
| 1 | 6 | 19.915 | 15.2 | 15.8 | 20.7 | 13.6732 | 12.64 | 14.15 | 8.15 | 13.1 | 13.78 | 22.9594 | 12.9 | 14.5 | 15.65 | 13.9 | 13.2 | 14 | 14.7 | 12.4 | 14.55 | 15.9 | 27.34 | 4.311 | 14.6202
| Mean |  | 20.0209 | 15.5167 | 15.5667 | 20.4667 | 13.8476 | 19.7117 | 14.4917 | 8.495 | 12.9833 | 13.8883 | 23.431 | 13.8 | 13.85 | 16.1917 | 13.975 | 13.6333 | 14.4 | 15.1833 | 11.9667 | 14.6 | 15.4333 | 27.94 | 4.207 | 14.6586
| Median |  | 20.0201 | 15.4 | 15.8 | 20.4 | 13.8457 | 20.72 | 14.55 | 8.5 | 13.05 | 13.82 | 23.4708 | 13.65 | 13.65 | 16.175 | 13.925 | 13.5 | 14.55 | 15.2 | 12.2 | 14.6 | 15.4 | 27.695 | 4.271 | 14.7792
| Std Dev. |  | 0.076248 | 0.33116 | 0.50859 | 0.41312 | 0.79177 | 3.6439 | 0.35555 | 0.28988 | 0.3656 | 0.39822 | 0.59476 | 1.0354 | 0.40373 | 0.42827 | 0.16355 | 0.75542 | 0.38341 | 0.4916 | 0.72296 | 0.59749 | 0.45019 | 1.0402 | 0.19978 | 0.66852
| Rel.Std.Dev. |  | 0.3808% | 2.134% | 3.267% | 2.018% | 5.718% | 18.49% | 2.453% | 3.412% | 2.816% | 2.867% | 2.538% | 7.503% | 2.915% | 2.645% | 1.17% | 5.541% | 2.663% | 3.238% | 6.041% | 4.092% | 2.917% | 3.723% | 4.749% | 4.561%
| PDM3 |  | 39.3% | 7.959% | 8.307% | 42.4% | -3.654% | 37.15% | 0.8276% | -40.89% | -9.667% | -3.37% | 63.02% | -3.985% | -3.637% | 12.66% | -2.767% | -5.144% | 0.1898% | 5.64% | -16.74% | 1.581% | 7.379% | 94.4% | -70.73% | 1.989%
| Z-Score (Absolute) |  | 3.92 | 0.7 | 0.73 | 4.24 | 0.5 | 3.7 | 0.04 | 4.33 | 1.12 | 0.47 | 6.36 | 0.53 | 0.5 | 1.18 | 0.41 | 0.65 | 0.1 | 0.46 | 1.85 | 0.04 | 0.64 | 9.59 | 7.4 | 0.08
| Analytical results for Re in OREAS 45f (Indicative Value < 0.001 ppm) | [31 empty] | Indicative
Re | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab; 25: Lab; 27: Lab |  | 1
| No | No | 02: 04; 04: 07; 05: 13; 07: 14; 10: 18; 13: 19; 14: 21; 15: 24; 17: 28; 18: 29; 19: 30; 20: 31; 21: 34; 23: 35; 24: 36; 25: 38; 27: 39 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*MS; 14: AR*OES; 15: AR*OES/MS; 17: AR*MS; 18: AR*MS; 19: AR*MS; 20: AR*MS; 21: AR*MS; 23: AR*MS; 24: AR*OES/MS; 25: AR*MS; 27: AR*MS |  | 3
|  |  | 02: 0.25g; 04: 01g; 05: 25g; 07: 0.5g; 10: 0.4g; 13: 40g; 14: 0.15g; 15: 01g; 17: 0.5g; 18: 0.5g; 19: 0.5g; 20: 0.5g; 21: 0.25g; 23: 0.5g; 24: 0.5g; 25: 01g; 27: 0.5g |  | 3
| 1 | 1 | 02: < 0.002; 04: < 0.001; 05: < 0.001; 07: < 0.001; 13: < 0.01; 14: < 5; 15: < 0.001; 17: < 0.001; 18: < 0.001; 19: < 0.001; 20: < 0.001; 21: < 0.1; 23: < 0.001; 24: 0.005; 25: < 0.001; 27: < 0.002 |  | 1
| 1 | 2 | 02: < 0.002; 04: < 0.001; 05: < 0.001; 07: < 0.001; 13: < 0.01; 14: < 5; 15: < 0.001; 17: < 0.001; 18: < 0.001; 19: < 0.001; 20: < 0.001; 21: < 0.1; 23: < 0.001; 24: 0.006; 25: < 0.001; 27: < 0.002 |  | 11
| 1 | 3 | 02: < 0.002; 04: < 0.001; 05: < 0.001; 07: < 0.001; 13: < 0.01; 14: < 5; 15: < 0.001; 17: < 0.001; 18: < 0.001; 19: < 0.001; 20: < 0.001; 21: < 0.1; 23: < 0.001; 24: 0.006; 25: < 0.001; 27: < 0.002 |  | 16
| 1 | 4 | 02: < 0.002; 04: < 0.001; 05: < 0.001; 07: < 0.001; 13: < 0.01; 14: < 5; 15: < 0.001; 17: < 0.001; 18: < 0.001; 19: < 0.001; 20: < 0.001; 21: < 0.1; 23: < 0.001; 24: 0.005; 25: < 0.001; 27: < 0.002 |  | < 0.001
| 1 | 5 | 02: < 0.002; 04: < 0.001; 05: < 0.001; 07: < 0.001; 13: < 0.01; 14: < 5; 15: < 0.001; 17: < 0.001; 18: < 0.001; 19: < 0.001; 20: < 0.001; 21: < 0.1; 23: < 0.001; 24: 0.005; 25: < 0.001; 27: < 0.002 |  | 29
| 1 | 6 | 02: 0.002; 04: < 0.001; 05: < 0.001; 07: < 0.001; 13: < 0.01; 14: < 5; 15: < 0.001; 17: < 0.001; 18: < 0.001; 19: < 0.001; 20: < 0.001; 21: < 0.1; 23: < 0.001; 24: 0.004; 25: < 0.001; 27: < 0.002
| Mean |  | 02: 0.002; 24: 0.0051667
| Median |  | 02: 0.002; 24: 0.005
| Std Dev. |  | 24: 0.00075277
| Rel.Std.Dev. |  | 24: 14.57%
PDM3
Z-Score (Absolute) | bar: 02 to 27
Results from laboratories 19, 21 & 34 were removed due to their 10 ppm or greater reading resolution.
| Analytical results for Rh in OREAS 45f (Indicative Value < 10 ppb) | [31 empty] | Indicative
Rh | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 39 |  | ppb
|  |  | 02: AR*MS |  | 1
|  |  | 02: 0.5g |  | 1
| 1 | 1 | 02: < 10 |  | 1
| 1 | 2 | 02: < 10 |  | 12
| 1 | 3 | 02: < 10 |  | 16
| 1 | 4 | 02: < 10 |  | < 10
| 1 | 5 | 02: < 10 |  | 30
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for Ru in OREAS 45f (Indicative Value < 10 ppb) | [31 empty] | Indicative
Ru | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 39 |  | ppb
|  |  | 02: AR*MS |  | 1
|  |  | 02: 0.5g |  | 1
| 1 | 1 | 02: < 10 |  | 1
| 1 | 2 | 02: < 10 |  | 13
| 1 | 3 | 02: < 10 |  | 16
| 1 | 4 | 02: < 10 |  | < 10
| 1 | 5 | 02: < 10 |  | 31
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for S in OREAS 45f (Certified Value 0.027 wt.%) | [31 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 14 | 15 | 17 | 18 | 20 | 21 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | [6 empty] | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | [6 empty] | 3
|  |  | 0.5g | 0.25g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | 0.4g | 15g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.027987 | 0.03 | 0.02 | 0.026 | 0.02 | 0.0326 | 0.03 | 0.026 | < 0.02 | < 0.05 | 0.02679 | 0.0271 | 0.01 | 0.03 | 0.03 | 0.02 | 0.0285 | 0.03 | 0.03 | 0.03 | 0.03 | 0.027567 | 0.022917 | 0.027924 | [6 empty] | 1
| 1 | 2 | 0.028007 | 0.02 | 0.02 | 0.026 | 0.02 | 0.0303 | 0.02 | 0.026 | < 0.02 | < 0.05 | 0.028809 | 0.0263 | 0.02 | 0.03 | 0.03 | 0.02 | 0.0275 | 0.02 | 0.03 | 0.02 | 0.03 | 0.0273 | 0.022556 | 0.028338 | [6 empty] | 28
| 1 | 3 | 0.028261 | 0.03 | 0.02 | 0.026 | 0.02 | 0.0354 | 0.03 | 0.026 | < 0.02 | < 0.05 | 0.027746 | 0.0275 | 0.01 | 0.03 | 0.03 | 0.02 | 0.027 | 0.02 | 0.03 | 0.02 | 0.03 | 0.027233 | 0.02367 | 0.02889 | [6 empty] | 16
| 1 | 4 | 0.028411 | 0.03 | 0.02 | 0.027 | 0.02 | 0.0335 | 0.03 | 0.0265 | < 0.02 | < 0.05 | 0.027177 | 0.0272 | 0.01 | 0.03 | 0.03 | 0.02 | 0.0265 | 0.02 | 0.03 | 0.03 | 0.03 | 0.027667 | 0.023097 | 0.027531 | [6 empty] | 0.026561
| 1 | 5 | 0.02821 | 0.03 | 0.02 | 0.027 | 0.02 | 0.0318 | 0.03 | 0.0255 | < 0.02 | < 0.05 | 0.025778 | 0.0275 | 0.01 | 0.03 | 0.03 | 0.02 | 0.028 | 0.03 | 0.03 | 0.02 | 0.03 | 0.027 | 0.023161 | 0.02835 | [6 empty] | 175
| 1 | 6 | 0.028188 | 0.02 | 0.02 | 0.027 | 0.02 | 0.0332 | 0.03 | 0.026 | < 0.02 | < 0.05 | 0.027538 | 0.0272 | 0.01 | 0.03 | 0.04 | 0.02 | 0.029 | 0.03 | 0.03 | 0.02 | 0.03 | 0.027333 | 0.02289 | 0.028278
| Mean |  | 0.028177 | 0.026667 | 0.02 | 0.0265 | 0.02 | 0.0328 | 0.028333 | 0.026 |  |  | 0.027306 | 0.027133 | 0.011667 | 0.03 | 0.031667 | 0.02 | 0.02775 | 0.025 | 0.03 | 0.023333 | 0.03 | 0.02735 | 0.023049 | 0.028219
| Median |  | 0.028199 | 0.03 | 0.02 | 0.0265 | 0.02 | 0.0329 | 0.03 | 0.026 |  |  | 0.027357 | 0.0272 | 0.01 | 0.03 | 0.03 | 0.02 | 0.02775 | 0.025 | 0.03 | 0.02 | 0.03 | 0.027317 | 0.023007 | 0.028308
| Std Dev. |  | 0.00016001 | 0.005164 | 0 | 0.00054772 | 0 | 0.0017146 | 0.0040825 | 0.00031623 |  |  | 0.0010122 | 0.00044121 | 0.0040825 | 0 | 0.0040825 | 0 | 0.00093541 | 0.0054772 | 0 | 0.005164 | 0 | 0.00023921 | 0.00037051 | 0.00045734
| Rel.Std.Dev. |  | 0.5679% | 19.36% | 0% | 2.067% | 0% | 5.228% | 14.41% | 1.216% |  |  | 3.707% | 1.626% | 34.99% | 0% | 12.89% | 0% | 3.371% | 21.91% | 0% | 22.13% | 0% | 0.8746% | 1.608% | 1.621%
| PDM3 |  | 6.085% | 0.3972% | -24.7% | -0.2302% | -24.7% | 23.49% | 6.672% | -2.113% |  |  | 2.805% | 2.154% | -56.08% | 12.95% | 19.22% | -24.7% | 4.476% | -5.878% | 12.95% | -12.15% | 12.95% | 2.97% | -13.22% | 6.24%
| Z-Score (Absolute) |  | 0.34 | 0.06 | 1.86 | 0.11 | 1.86 | 1.59 | 0.39 | 0.24 | 4.56 | 0.51 | 0.11 | 0.06 | 4.11 | 0.84 | 1.29 | 1.86 | 0.23 | 0.51 | 0.84 | 0.96 | 0.84 | 0.12 | 1.04 | 0.36
| Analytical results for Sb in OREAS 45f (Indicative Value 0.24 ppm) | [31 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 0.8335 | 0.33 | 0.31 | 0.33 | 0.79553 | 0.15 | 0.193 | 1.45 | < 5 | 0.2 | 0.31 | 2.1 | 0.2 | 0.35575 | 0.259 | 0.16 | 0.215 | 0.195 | 0.17 | 0.224 | 0.3 | 0.41 | 0.223 | 0.66 | < 5 | 0.24667 | 0.061 | 0.27439 |  |  | 1
| 1 | 2 | 0.8576 | 0.28 | 0.31 | 0.33 | 0.8763 | 0.16 | 0.169 | 1.49 | < 5 | 0.2 | 0.54 | 2.5 | 0.18 | 0.3531 | 0.258 | 0.17 | 0.165 | 0.197 | 0.15 | 0.229 | 0.2 | 0.43 | 0.267 | 0.65 | < 5 | 0.24667 | < 0.05 | 0.28749 |  |  | 26
| 1 | 3 | 0.83006 | 0.32 | 0.33 | 0.29 | 0.81305 | 0.16 | 0.208 | 1.6 | < 5 | 0.2 | 0.32 | 1.6 | 0.19 | 0.3696 | 0.356 | 0.15 | 0.276 | 0.18 | 0.13 | 0.263 | 0.3 | 0.48 | 0.266 | 0.58 | < 5 | 0.19333 | 0.056 | 0.30961 |  |  | 16
| 1 | 4 | 0.8496 | 0.26 | 0.32 | 0.31 | 0.77036 | 0.16 | 0.181 | 1.58 | < 5 | 0.2 | 0.39 | 2 | 0.21 | 0.371 | 0.273 | 0.14 | 0.226 | 0.186 | 0.14 | 0.249 | 0.3 | 0.4 | 0.282 | 0.57 | < 5 | 0.26333 | 0.051 | 0.28765 |  |  | 0.24367
| 1 | 5 | 0.8589 | 0.33 | 0.32 | 0.28 | 0.83018 | 0.15 | 0.174 | 1.49 | < 5 | 0.2 | 0.33 | 2 | 0.19 | 0.35332 | 0.296 | 0.16 | 0.204 | 0.2 | 0.17 | 0.244 | 0.3 | 0.35 | 0.286 | 0.61 | < 5 | 0.25 | 0.058 | 0.28035 |  |  | 32
| 1 | 6 | 0.853 | 0.32 | 0.3 | 0.29 | 0.79031 | 0.16 | 0.184 | 1.59 | < 5 | 0.1 | 0.27 | 2.2 | 0.17 | 0.35201 | 0.236 | 0.16 | 0.185 | 0.183 | 0.16 | 0.249 | 0.2 | 0.38 | 0.296 | 0.56 | < 5 | 0.24 | 0.054 | 0.27245
| Mean |  | 0.84711 | 0.30667 | 0.315 | 0.305 | 0.81262 | 0.15667 | 0.18483 | 1.5333 |  | 0.18333 | 0.36 | 2.0667 | 0.19 | 0.35913 | 0.27967 | 0.15667 | 0.21183 | 0.19017 | 0.15333 | 0.243 | 0.26667 | 0.40833 | 0.27 | 0.605 |  | 0.24 | 0.056 | 0.28532
| Median |  | 0.8513 | 0.32 | 0.315 | 0.3 | 0.80429 | 0.16 | 0.1825 | 1.535 |  | 0.2 | 0.325 | 2.05 | 0.19 | 0.35453 | 0.266 | 0.16 | 0.2095 | 0.1905 | 0.155 | 0.2465 | 0.3 | 0.405 | 0.2745 | 0.595 |  | 0.24667 | 0.056 | 0.28392
| Std Dev. |  | 0.012376 | 0.029439 | 0.010488 | 0.021679 | 0.037254 | 0.005164 | 0.014049 | 0.064083 |  | 0.040825 | 0.096333 | 0.29439 | 0.014142 | 0.0087495 | 0.042288 | 0.010328 | 0.038217 | 0.0082321 | 0.01633 | 0.014353 | 0.05164 | 0.04446 | 0.025729 | 0.042308 |  | 0.024129 | 0.0038079 | 0.01349
| Rel.Std.Dev. |  | 1.461% | 9.6% | 3.33% | 7.108% | 4.584% | 3.296% | 7.601% | 4.179% |  | 22.27% | 26.76% | 14.24% | 7.443% | 2.436% | 15.12% | 6.592% | 18.04% | 4.329% | 10.65% | 5.906% | 19.36% | 10.89% | 9.529% | 6.993% |  | 10.05% | 6.8% | 4.728%
| PDM3 |  | 247.6% | 25.85% | 29.27% | 25.17% | 233.5% | -35.7% | -24.15% | 529.3% |  | -24.76% | 47.74% | 748.1% | -22.02% | 47.38% | 14.77% | -35.7% | -13.06% | -21.96% | -37.07% | -0.274% | 9.439% | 67.58% | 10.81% | 148.3% |  | -1.505% | -77.02% | 17.1%
| Z-Score (Absolute) |  | 3.45 | 0.07 | 0.12 | 0.06 | 3.24 | 0.87 | 0.69 | 7.75 | 13.8 | NA | 0.41 | 11.09 | 0.66 | 0.4 | 0.1 | 0.87 | 0.52 | 0.66 | 0.89 | 0.33 | NA | 0.71 | 0.16 | 1.94 | 13.8 | 0.35 | 1.53 | 0.06
Results from laboratory 19 & 33 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Sc in OREAS 45f (Certified Value 31.4 ppm) | [31 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 2
| 1 | 1 | 31.0212 | 30.3 | 29.4 | 31 | 61.1489 | 36.2 | 31.1 | 34.6 | 31 | 22 | 28 | 29.4 | 29 | 30.43 | 30.2366 | 29 | 38.8 | 31.2 | 29.8 | 34.7 | 34 | 33.4 | 33 | 30.4 | 27.9 | 32.3 | 35.6667 | 14.35 | 31.8312 |  | 1
| 1 | 2 | 31.0856 | 29.5 | 29.5 | 31 | 53.6871 | 34 | 29.6 | 32.8 | 31 | 21.2 | 26 | 30.5 | 30 | 30.77 | 30.7203 | 28.6 | 35.3 | 32.5 | 30.8 | 32.9 | 34.1 | 27.1 | 32.9 | 29.9 | 33.3 | 31.7 | 34.1333 | 13.99 | 30.8651 |  | #N/A
| 1 | 3 | 31.0728 | 29.6 | 29.8 | 30 | 57.0976 | 35.8 | 30.3 | 33.7 | 30 | 22.2 | 28 | 28 | 30 | 30.2 | 31.2432 | 27.7 | 33.4 | 34.4 | 30.4 | 33.2 | 33.8 | 27.3 | 30.4 | 33.8 | 35.1 | 32.5 | 33.5333 | 14.32 | 31.565 |  | 16
| 1 | 4 | 31.056 | 29.9 | 30.5 | 31 | 55.6979 | 35.5 | 30.4 | 33.7 | 31 | 21.8 | 28 | 27.8 | 30 | 28.61 | 31.4359 | 29.2 | 34 | 33.8 | 29.8 | 34.5 | 34.1 | 31.2 | 31.7 | 34.3 | 33.1 | 31.3 | 34.7667 | 14.48 | 31.2861 |  | 31.4172
| 1 | 5 | 31.1452 | 28.3 | 30 | 31 | 60.5442 | 34.9 | 29.4 | 33.3 | 30 | 21.4 | 26 | 29.5 | 30 | 28.53 | 30.219 | 28.3 | 34.5 | 33.4 | 29.5 | 33.9 | 34.2 | 28.9 | 32.6 | 32.7 | 32.6 | 31.9 | 34.6 | 14.27 | 31.7476 |  | 176
| 1 | 6 | 31.196 | 28.8 | 29.5 | 31 | 55.6008 | 32.1 | 29.6 | 34.3 | 31 | 19.9 | 28 | 29.1 | 30 | 30.12 | 31.1528 | 27.4 | 37.2 | 32 | 28.4 | 34 | 33.5 | 29.4 | 32.6 | 32.8 | 33.3 | 31.2 | 34.3667 | 14.42 | 32.0542
| Mean |  | 31.0961 | 29.4 | 29.7833 | 30.8333 | 57.2961 | 34.75 | 30.0667 | 33.7333 | 30.6667 | 21.4167 | 27.3333 | 29.05 | 29.8333 | 29.7767 | 30.8346 | 28.3667 | 35.5333 | 32.8833 | 29.7833 | 33.8667 | 33.95 | 29.55 | 32.2 | 32.3167 | 32.55 | 31.8167 | 34.5111 | 14.305 | 31.5582
| Median |  | 31.0792 | 29.55 | 29.65 | 31 | 56.3978 | 35.2 | 29.95 | 33.7 | 31 | 21.6 | 28 | 29.25 | 30 | 30.16 | 30.9365 | 28.45 | 34.9 | 32.95 | 29.8 | 33.95 | 34.05 | 29.15 | 32.6 | 32.75 | 33.2 | 31.8 | 34.4833 | 14.335 | 31.6563
| Std Dev. |  | 0.063658 | 0.73212 | 0.41673 | 0.40825 | 2.9627 | 1.5083 | 0.65013 | 0.6532 | 0.5164 | 0.83046 | 1.0328 | 1.0095 | 0.40825 | 0.96187 | 0.52523 | 0.71181 | 2.0724 | 1.1974 | 0.82563 | 0.70616 | 0.25884 | 2.4106 | 0.99398 | 1.7904 | 2.4313 | 0.52313 | 0.71108 | 0.1712 | 0.42682
| Rel.Std.Dev. |  | 0.2047% | 2.49% | 1.399% | 1.324% | 5.171% | 4.34% | 2.162% | 1.936% | 1.684% | 3.878% | 3.779% | 3.475% | 1.368% | 3.23% | 1.703% | 2.509% | 5.832% | 3.641% | 2.772% | 2.085% | 0.7624% | 8.158% | 3.087% | 5.54% | 7.469% | 1.644% | 2.06% | 1.197% | 1.353%
| PDM3 |  | -1.022% | -6.421% | -5.2% | -1.858% | 82.37% | 10.61% | -4.299% | 7.372% | -2.389% | -31.83% | -13% | -7.535% | -5.041% | -5.222% | -1.854% | -9.71% | 13.1% | 4.667% | -5.2% | 7.797% | 8.062% | -5.943% | 2.492% | 2.863% | 3.606% | 1.272% | 9.848% | -54.47% | 0.4489%
| Z-Score (Absolute) |  | 0.12 | 0.65 | 0.48 | 0 | 12.04 | 1.78 | 0.35 | 1.32 | 0.08 | 4.29 | 1.59 | 0.81 | 0.46 | 0.48 | 0 | 1.12 | 2.14 | 0.93 | 0.48 | 1.38 | 1.42 | 0.58 | 0.62 | 0.67 | 0.78 | 0.45 | 1.67 | 7.52 | 0.33
| Analytical results for Se in OREAS 45f (Indicative Value 0.58 ppm) | [31 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 0.5g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | < 1 | 1 | 2.359 | < 0.5 | 0.5 |  | < 1 | < 0.5 | < 0.5 | < 5 | 0.5 | < 1 | < 1 | < 1 | 0.6 | 0.7 | 0.3 | 0.6 | < 1 | < 1 | 0.6 | 1.6 | 1 | < 0.2 | 0.52364 |  |  |  |  | 1
| 1 | 2 | < 1 | < 1 | 1 | 2.3174 | 0.5 | 0.4 |  | < 1 | < 0.5 | < 0.5 | < 5 | 0.5 | 1.1535 | < 1 | < 1 | 0.6 | 0.7 | 0.2 | 0.6 | < 1 | < 1 | 0.5 | 1.2 | 1 | < 0.2 | 0.51393 |  |  |  |  | 8
| 1 | 3 | < 1 | < 1 | 1 | 2.1852 | < 0.5 | 0.5 |  | < 1 | < 0.5 | < 0.5 | < 5 | 0.6 | 1.0501 | < 1 | < 1 | 0.7 | 0.6 | 0.4 | 0.7 | < 1 | < 1 | 0.7 | 1.3 | 1 | < 0.2 | 0.54281 |  |  |  |  | 16
| 1 | 4 | < 1 | < 1 | 1 | 3.3343 | < 0.5 | 0.5 |  | < 1 | < 0.5 | 0.6 | < 5 | 0.6 | < 1 | < 1 | < 1 | 0.6 | 0.6 | 0.5 | 0.7 | < 1 | < 1 | 0.6 | 1.4 | 1 | < 0.2 | 0.56883 |  |  |  |  | 0.57696
| 1 | 5 | < 1 | < 1 | 1 | 2.1343 | < 0.5 | 0.5 |  | < 1 | < 0.5 | < 0.5 | < 5 | 0.6 | 1.0363 | < 1 | < 1 | 0.6 | 0.6 | 0.3 | 0.6 | < 1 | < 1 | 0.7 | 1.4 | 1 | < 0.2 | 0.5553 |  |  |  |  | 27
| 1 | 6 | < 1 | < 1 | 1 | 2.1674 | < 0.5 | 0.4 |  | < 1 | < 0.5 | < 0.5 | < 5 | 0.5 | 1.1985 | < 1 | < 1 | 0.6 | 0.7 | 0.2 | 0.6 | < 1 | < 1 | 0.7 | 1.1 | 1 | < 0.2 | 0.55157
| Mean |  | 05: 1; 07: 2.4162; 10: 0.5; 13: 0.46667; 18: 0.6; 20: 0.55; 21: 1.1096; 25: 0.61667; 27: 0.65; 28: 0.31667; 29: 0.63333; 33: 0.63333; 34: 1.3333; 35: 1; 37: 0.54268
| Median |  | 05: 1; 07: 2.2513; 10: 0.5; 13: 0.5; 18: 0.6; 20: 0.55; 21: 1.1018; 25: 0.6; 27: 0.65; 28: 0.3; 29: 0.6; 33: 0.65; 34: 1.35; 35: 1; 37: 0.54719
| Std Dev. |  | 05: 0; 07: 0.45842; 13: 0.05164; 20: 0.054772; 21: 0.079068; 25: 0.040825; 27: 0.054772; 28: 0.1169; 29: 0.05164; 33: 0.08165; 34: 0.17512; 35: 0; 37: 0.020549
| Rel.Std.Dev. |  | 05: 0%; 07: 18.97%; 13: 11.07%; 20: 9.959%; 21: 7.126%; 25: 6.62%; 27: 8.427%; 28: 36.92%; 29: 8.154%; 33: 12.89%; 34: 13.13%; 35: 0%; 37: 3.786%
| PDM3 |  | 05: 73.32%; 07: 318.8%; 10: -13.34%; 13: -19.12%; 18: 3.993%; 20: -4.674%; 21: 92.32%; 25: 6.881%; 27: 12.66%; 28: -45.12%; 29: 9.77%; 33: 9.77%; 34: 131.1%; 35: 73.32%; 37: -5.942%
| Z-Score (Absolute) |  | 0 | 0 | NA | 9.69 | 1.05 | 0.17 | NA | 0 | 1.26 | 0.97 | 10.11 | 0.25 | 2.06 | 0 | 0 | 0.59 | 0.76 | 0.93 | 0.67 | 0 | 0 | 0.67 | 4.21 | NA | 2.02 | 0.22
Results from laboratory 7 & 38 were removed due to their 1 ppm or greater reading resolution.
| Analytical results for Sm in OREAS 45f (Certified Value 1.91 ppm) | [31 empty] | Certified
Sm | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 15: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 18; 10: 23; 13: 24; 14: 33; 15: 38; 17: 39; 18: 40 |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*MS; 17: AR*OES; 18: AR*MS |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.4g; 10: 0.5g; 13: 01g; 14: 0.4g; 15: 01g; 17: 0.5g; 18: 0.5g |  | 2
| 1 | 1 | 02: 2.109; 04: 2.07; 05: 1.8; 07: 1.55; 10: 1.65; 13: 2.0918; 14: 1.89; 15: 2.39; 17: 1.013; 18: 1.7428 |  | 1
| 1 | 2 | 02: 2.1331; 04: 2.15; 05: 1.89; 07: 1.56; 10: 1.77; 13: 2.0402; 14: 1.54; 15: 2.2533; 17: 0.8; 18: 1.8179 |  | #N/A
| 1 | 3 | 02: 2.1503; 04: 2.03; 05: 1.89; 07: 1.59; 10: 1.66; 13: 2.0774; 14: 1.46; 15: 2.3367; 17: 0.882; 18: 1.7432 |  | 16
| 1 | 4 | 02: 2.1325; 04: 2.06; 05: 1.95; 07: 1.6; 10: 1.61; 13: 2.1251; 14: 1.72; 15: 2.3867; 17: 0.757; 18: 1.8115 |  | 1.9094
| 1 | 5 | 02: 2.0607; 04: 2; 05: 1.79; 07: 1.56; 10: 1.69; 13: 2.0822; 14: 1.6; 15: 2.31; 17: 0.976; 18: 1.8495 |  | 177
| 1 | 6 | 02: 2.1095; 04: 2.16; 05: 1.66; 07: 1.46; 10: 1.69; 13: 2.0118; 14: 1.83; 15: 2.3167; 17: 0.993; 18: 1.778
| Mean |  | 02: 2.1159; 04: 2.0783; 05: 1.83; 07: 1.5533; 10: 1.6783; 13: 2.0714; 14: 1.6733; 15: 2.3322; 17: 0.9035; 18: 1.7905
| Median |  | 02: 2.121; 04: 2.065; 05: 1.845; 07: 1.56; 10: 1.675; 13: 2.0798; 14: 1.66; 15: 2.3267; 17: 0.929; 18: 1.7948
| Std Dev. |  | 02: 0.031274; 04: 0.064317; 05: 0.10296; 07: 0.049666; 10: 0.053821; 13: 0.039978; 14: 0.16872; 15: 0.051539; 17: 0.1076; 18: 0.043227
| Rel.Std.Dev. |  | 02: 1.478%; 04: 3.095%; 05: 5.626%; 07: 3.197%; 10: 3.207%; 13: 1.93%; 14: 10.08%; 15: 2.21%; 17: 11.91%; 18: 2.414%
| PDM3 |  | 02: 10.81%; 04: 8.847%; 05: -4.159%; 07: -18.65%; 10: -12.1%; 13: 8.485%; 14: -12.36%; 15: 22.14%; 17: -52.68%; 18: -6.228%
| Z-Score (Absolute) |  | 02: 0.8; 04: 0.7; 05: 0.05; 07: 0.67; 10: 0.34; 13: 0.68; 14: 0.36; 15: 1.36; 17: 2.36; 18: 0.05
| Analytical results for Sn in OREAS 45f (Certified Value 1.97 ppm) | [31 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 2.3061 | 2.2 | 1.6 | 1.8 | 4.5428 | 2.1 | 2.07 | < 5 | 1.8 | < 2 | < 10 | 1.7 | 2.0315 | 1.84 | 1.4 | 1.93 | 1.97 | 2 | 1.97 | 2 | 2.3 | 1.85 | 1.2 | < 10 | 2.3933 | 0.746 | 1.8266 |  |  |  | 1
| 1 | 2 | 2.2998 | 2 | 1.7 | 1.8 | 4.9844 | 2 | 2.09 | < 5 | 1.7 | < 2 | < 10 | 1.8 | 2.0682 | 1.79 | 1.4 | 2.03 | 1.94 | 1.9 | 1.93 | 1.8 | 2.3 | 1.75 | 1 | < 10 | 2.3267 | 0.695 | 1.8086 |  |  |  | 3
| 1 | 3 | 2.3198 | 2.1 | 1.7 | 1.8 | 4.8501 | 2.1 | 2.08 | < 5 | 1.8 | < 2 | < 10 | 1.7 | 2.1815 | 1.93 | 1.3 | 2 | 2.03 | 1.8 | 1.89 | 1.9 | 2.3 | 1.95 | 0.9 | < 10 | 2.24 | 0.77 | 1.8723 |  |  |  | 16
| 1 | 4 | 2.3106 | 2 | 1.6 | 1.7 | 4.342 | 2.2 | 2.07 | < 5 | 1.8 | < 2 | < 10 | 1.7 | 2.1702 | 1.9 | 1.3 | 1.91 | 1.99 | 2 | 1.96 | 1.7 | 2.2 | 1.93 | 0.9 | < 10 | 2.3133 | 0.713 | 1.862 |  |  |  | 1.9655
| 1 | 5 | 2.2935 | 2.4 | 1.6 | 1.6 | 4.1483 | 2 | 2.07 | < 5 | 1.6 | < 2 | < 10 | 1.7 | 2.0768 | 1.88 | 1.5 | 2.01 | 2.01 | 1.9 | 1.94 | 2 | 2.3 | 1.88 | 1 | < 10 | 2.3267 | 0.739 | 1.8245 |  |  |  | 178
| 1 | 6 | 2.3405 | 2 | 1.6 | 1.8 | 4.4572 | 1.8 | 2.07 | < 5 | 1.6 | < 2 | < 10 | 1.8 | 2.0305 | 1.81 | 1.5 | 2.03 | 1.95 | 1.9 | 1.89 | 1.8 | 2.3 | 1.95 | 1.1 | < 10 | 2.3033 | 0.744 | 1.7672
| Mean |  | 2.3117 | 2.1167 | 1.6333 | 1.75 | 4.5541 | 2.0333 | 2.075 |  | 1.7167 |  |  | 1.7333 | 2.0931 | 1.8583 | 1.4 | 1.985 | 1.9817 | 1.9167 | 1.93 | 1.8667 | 2.2833 | 1.885 | 1.0167 |  | 2.3172 | 0.7345 | 1.8269
| Median |  | 2.3083 | 2.05 | 1.6 | 1.8 | 4.5 | 2.05 | 2.07 |  | 1.75 |  |  | 1.7 | 2.0725 | 1.86 | 1.4 | 2.005 | 1.98 | 1.9 | 1.935 | 1.85 | 2.3 | 1.905 | 1 |  | 2.32 | 0.7415 | 1.8255
| Std Dev. |  | 0.016736 | 0.16021 | 0.05164 | 0.083666 | 0.31368 | 0.13663 | 0.0083666 |  | 0.098319 |  |  | 0.05164 | 0.066874 | 0.054191 | 0.089443 | 0.052058 | 0.034881 | 0.075277 | 0.034059 | 0.12111 | 0.040825 | 0.077395 | 0.1169 |  | 0.04928 | 0.026554 | 0.037947
| Rel.Std.Dev. |  | 0.724% | 7.569% | 3.162% | 4.781% | 6.888% | 6.719% | 0.4032% |  | 5.727% |  |  | 2.979% | 3.195% | 2.916% | 6.389% | 2.623% | 1.76% | 3.928% | 1.765% | 6.488% | 1.788% | 4.106% | 11.5% |  | 2.127% | 3.615% | 2.077%
| PDM3 |  | 17.62% | 7.693% | -16.9% | -10.96% | 131.7% | 3.453% | 5.573% |  | -12.66% |  |  | -11.81% | 6.495% | -5.451% | -28.77% | 0.9937% | 0.8241% | -2.483% | -1.805% | -5.027% | 16.17% | -4.094% | -48.27% |  | 17.9% | -62.63% | -7.052%
| Z-Score (Absolute) |  | 1.31 | 0.64 | 1.02 | 0.62 | 9 | 0.35 | 0.5 | 1.95 | 0.73 | 3.19 | 10.53 | 0.67 | 0.56 | 0.25 | 1.82 | 0.19 | 0.18 | 0.05 | 0 | 0.22 | 1.21 | 0.15 | 3.13 | 10.53 | 1.33 | 4.1 | 0.35
| Analytical results for Sr in OREAS 45f (Certified Value 13.2 ppm) | [31 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES |  |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 14.788 | 13.3 | 12.7 | 13.2 | 18.8193 | 14.5 | 13 | 48 | 22 | 12 | 14.3 | 12 | 12.66 | 15.4871 | 11.4 | 8.3 | 13.15 | 13 | 14.1 | 14.25 | 12.2 | 11.1 | 12.5 | 16.1 | 12.1 | 19.0433 | 9.633 | 12.803 |  |  | 1
| 1 | 2 | 14.813 | 13 | 13.1 | 14.1 | 18.5721 | 13.8 | 12.95 | 47 | 22.2 | 11 | 14.4 | 12 | 13.13 | 15.2496 | 11.1 | 8.2 | 13.9 | 13.1 | 14 | 13.8 | 12.8 | 13 | 12.75 | 12.9 | 11.9 | 17.8467 | 9.738 | 12.7647 |  |  | #N/A
| 1 | 3 | 14.541 | 13.6 | 13.9 | 13.3 | 18.0326 | 14.3 | 13 | 47 | 22 | 11 | 14 | 12 | 12.64 | 15.4333 | 11.9 | 7.4 | 14.15 | 13.15 | 13.6 | 13.85 | 12.8 | 12.8 | 13.95 | 13 | 13 | 17.5933 | 9.755 | 13.1096 |  |  | 16
| 1 | 4 | 14.986 | 13.1 | 13.8 | 13.8 | 17.1724 | 15.2 | 12.85 | 48 | 21.1 | 11 | 14 | 12 | 12.19 | 15.9222 | 11.6 | 7.1 | 13.45 | 13.2 | 14.3 | 13.5 | 12.4 | 12.8 | 14.95 | 12.9 | 12.1 | 18.5067 | 9.765 | 13.1158 |  |  | 13.1565
| 1 | 5 | 14.055 | 13.7 | 13.5 | 13.4 | 18.6819 | 13.5 | 12.25 | 47 | 23.7 | 11 | 13.9 | 12 | 12.42 | 15.7086 | 11.8 | 7 | 14 | 13 | 14.1 | 13.45 | 12.2 | 11 | 13.7 | 12.1 | 12.7 | 18.05 | 9.861 | 13.3259 |  |  | 179
| 1 | 6 | 14.157 | 13.5 | 13.4 | 13.8 | 17.5398 | 11 | 12.4 | 47 | 19.9 | 11 | 14.3 | 12 | 12.62 | 15.7204 | 11.2 | 7.4 | 13.25 | 13 | 14.4 | 13.4 | 12.9 | 11 | 13.55 | 12.7 | 12.1 | 17.8967 | 9.695 | 13.0105
| Mean |  | 14.5567 | 13.3667 | 13.4 | 13.6 | 18.1364 | 13.7167 | 12.7417 | 47.3333 | 21.8167 | 11.1667 | 14.15 | 12 | 12.61 | 15.5869 | 11.5 | 7.5667 | 13.65 | 13.075 | 14.0833 | 13.7083 | 12.55 | 11.95 | 13.5667 | 13.2833 | 12.3167 | 18.1561 | 9.7412 | 13.0216
| Median |  | 14.6645 | 13.4 | 13.45 | 13.6 | 18.3024 | 14.05 | 12.9 | 47 | 22 | 11 | 14.15 | 12 | 12.63 | 15.5979 | 11.5 | 7.4 | 13.675 | 13.05 | 14.1 | 13.65 | 12.6 | 11.95 | 13.625 | 12.9 | 12.1 | 17.9733 | 9.7465 | 13.0601
| Std Dev. |  | 0.37821 | 0.28048 | 0.44721 | 0.35214 | 0.67076 | 1.4552 | 0.33078 | 0.5164 | 1.2608 | 0.40825 | 0.20736 | 0 | 0.3119 | 0.24182 | 0.32249 | 0.55377 | 0.42071 | 0.088034 | 0.27869 | 0.32468 | 0.32094 | 1.0075 | 0.88129 | 1.4176 | 0.43089 | 0.52936 | 0.076106 | 0.21125
| Rel.Std.Dev. |  | 2.598% | 2.098% | 3.337% | 2.589% | 3.698% | 10.61% | 2.596% | 1.091% | 5.779% | 3.656% | 1.465% | 0% | 2.473% | 1.551% | 2.804% | 7.319% | 3.082% | 0.6733% | 1.979% | 2.368% | 2.557% | 8.431% | 6.496% | 10.67% | 3.498% | 2.916% | 0.7813% | 1.622%
| PDM3 |  | 10.64% | 1.597% | 1.851% | 3.371% | 37.85% | 4.258% | -3.153% | 259.8% | 65.82% | -15.12% | 7.551% | -8.79% | -4.154% | 18.47% | -12.59% | -42.49% | 3.751% | -0.6196% | 7.045% | 4.194% | -4.61% | -9.171% | 3.117% | 0.9639% | -6.384% | 38% | -25.96% | -1.026%
| Z-Score (Absolute) |  | 0.98 | 0.01 | 0.01 | 0.18 | 3.99 | 0.28 | 0.54 | 28.5 | 7.08 | 1.86 | 0.64 | 1.16 | 0.65 | 1.85 | 1.58 | 4.88 | 0.22 | 0.26 | 0.59 | 0.27 | 0.7 | 1.2 | 0.15 | 0.08 | 0.9 | 4.01 | 3.06 | 0.3
| Analytical results for Ta in OREAS 45f (Indicative Value < 0.01 ppm) | [31 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [8 empty] | 3
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | [8 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | < 0.005 | 5 | < 0.05 | < 0.01 | < 0.05 | 0.06 | < 0.005 | < 0.005 | < 0.01 | < 0.005 | < 0.05 | < 0.05 | < 0.005 | < 0.01 | < 0.01 | < 0.02 | 0.00048 | [8 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | < 0.005 | < 5 | < 0.05 | < 0.01 | < 0.05 | 0.06 | < 0.005 | < 0.005 | < 0.01 | < 0.005 | < 0.05 | < 0.05 | < 0.005 | < 0.01 | < 0.01 | < 0.02 | 0.00055 | [8 empty] | 22
| 1 | 3 | < 0.05 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | 0.006 | < 5 | < 0.05 | < 0.01 | < 0.05 | 0.06 | < 0.005 | < 0.005 | < 0.01 | < 0.005 | < 0.05 | < 0.05 | < 0.005 | < 0.01 | < 0.01 | < 0.02 | 0.00056 | [8 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | < 0.005 | 5 | < 0.05 | < 0.01 | < 0.05 | 0.05 | < 0.005 | < 0.005 | < 0.01 | < 0.005 | < 0.05 | < 0.05 | < 0.005 | < 0.01 | < 0.01 | < 0.02 | 0.00034 | [8 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | 0.005 | 6 | < 0.05 | < 0.01 | < 0.05 | 0.06 | < 0.005 | < 0.005 | < 0.01 | < 0.005 | < 0.05 | < 0.05 | < 0.005 | < 0.01 | < 0.01 | < 0.02 | 0.00037 | [8 empty] | 28
| 1 | 6 | < 0.05 | < 0.05 | < 0.01 | < 0.1 | < 0.01 | 0.005 | 5 | < 0.05 | < 0.01 | < 0.05 | 0.06 | < 0.005 | < 0.005 | < 0.01 | < 0.005 | < 0.05 | < 0.05 | < 0.005 | < 0.01 | < 0.01 | < 0.02 | 0.00038
| Mean |  | 13: 0.0053333; 14: 5.25; 20: 0.058333; 34: 0.00044667
| Median |  | 13: 0.005; 14: 5; 20: 0.06; 34: 0.00043
| Std Dev. |  | 13: 0.00057735; 14: 0.5; 20: 0.0040825; 34: 9.6264e-05
| Rel.Std.Dev. |  | 13: 10.83%; 14: 9.524%; 20: 6.999%; 34: 21.55%
PDM3
| Z-Score (Absolute) |  | 3.82 | 3.82 | 0 | 8.6 | 0 | 0.21 | 827.59 | 3.82 | 0 | 3.82 | 10.2 | 0.48 | 0.48 | 0 | 0.48 | 3.82 | 3.82 | 0.48 | 0 | 0 | 0.96 | 0.87
| Analytical results for Tb in OREAS 45f (Certified Value 0.25 ppm) | [31 empty] | Certified
Tb | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 17: Lab; 18: Lab; 19: Lab; 20: Lab; 21: Lab; 23: Lab; 24: Lab; 25: Lab |  |  | 1
| No | No | 02: 02; 04: 04; 05: 05; 07: 07; 10: 13; 13: 18; 14: 23; 17: 24; 18: 25; 19: 27; 20: 33; 21: 34; 23: 38; 24: 39; 25: 40 |  |  | ppm
|  |  | 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 07: AR*OES/MS; 10: AR*OES/MS; 13: AR*MS; 14: AR*MS; 17: AR*OES/MS; 18: AR*MS; 19: AR*OES/MS; 20: AR*MS; 21: AR*MS; 23: AR*MS; 24: AR*OES; 25: AR*MS |  |  | 2
|  |  | 02: 0.5g; 04: 0.25g; 05: 0.25g; 07: 01g; 10: 25g; 13: 0.4g; 14: 0.5g; 17: 01g; 18: 0.25g; 19: 0.25g; 20: 0.4g; 21: 0.25g; 23: 01g; 24: 0.5g; 25: 0.5g |  |  | 3
| 1 | 1 | 02: < 0.5; 04: 0.27; 05: 0.27; 07: 0.255; 10: 0.25; 13: 0.24; 14: 0.236; 17: 0.24296; 18: 0.245; 19: 0.19; 20: 0.31; 21: 0.24; 23: 0.31333; 24: 0.52; 25: 0.23295 |  |  | 1
| 1 | 2 | 02: < 0.5; 04: 0.27; 05: 0.28; 07: 0.263; 10: 0.25; 13: 0.24; 14: 0.246; 17: 0.23918; 18: 0.232; 19: 0.19; 20: 0.27; 21: 0.24; 23: 0.3; 24: 0.57; 25: 0.24184 |  |  | #N/A
| 1 | 3 | 02: < 0.5; 04: 0.25; 05: 0.29; 07: 0.24; 10: 0.26; 13: 0.25; 14: 0.243; 17: 0.25171; 18: 0.253; 19: 0.2; 20: 0.25; 21: 0.22; 23: 0.30667; 24: 0.52; 25: 0.24076 |  |  | 16
| 1 | 4 | 02: < 0.5; 04: 0.25; 05: 0.28; 07: 0.257; 10: 0.25; 13: 0.24; 14: 0.231; 17: 0.25548; 18: 0.248; 19: 0.19; 20: 0.31; 21: 0.23; 23: 0.31; 24: 0.61; 25: 0.23933 |  |  | 0.24841
| 1 | 5 | 02: < 0.5; 04: 0.25; 05: 0.28; 07: 0.249; 10: 0.24; 13: 0.24; 14: 0.238; 17: 0.23557; 18: 0.24; 19: 0.2; 20: 0.26; 21: 0.24; 23: 0.30667; 24: 0.56; 25: 0.243 |  |  | 180
| 1 | 6 | 02: < 0.5; 04: 0.25; 05: 0.28; 07: 0.257; 10: 0.22; 13: 0.22; 14: 0.241; 17: 0.2473; 18: 0.247; 19: 0.19; 20: 0.3; 21: 0.24; 23: 0.30333; 24: 0.48; 25: 0.24449
| Mean |  | 04: 0.25667; 05: 0.28; 07: 0.2535; 10: 0.245; 13: 0.23833; 14: 0.23917; 17: 0.24537; 18: 0.24417; 19: 0.19333; 20: 0.28333; 21: 0.235; 23: 0.30667; 24: 0.54333; 25: 0.24039
| Median |  | 04: 0.25; 05: 0.28; 07: 0.256; 10: 0.25; 13: 0.24; 14: 0.2395; 17: 0.24513; 18: 0.246; 19: 0.19; 20: 0.285; 21: 0.24; 23: 0.30667; 24: 0.54; 25: 0.2413
| Std Dev. |  | 04: 0.010328; 05: 0.0063246; 07: 0.0079937; 10: 0.013784; 13: 0.0098319; 14: 0.0053448; 17: 0.0075677; 18: 0.0073052; 19: 0.005164; 20: 0.026583; 21: 0.0083666; 23: 0.004714; 24: 0.045898; 25: 0.004058
| Rel.Std.Dev. |  | 04: 4.024%; 05: 2.259%; 07: 3.153%; 10: 5.626%; 13: 4.125%; 14: 2.235%; 17: 3.084%; 18: 2.992%; 19: 2.671%; 20: 9.382%; 21: 3.56%; 23: 1.537%; 24: 8.448%; 25: 1.688%
| PDM3 |  | 04: 3.325%; 05: 12.72%; 07: 2.05%; 10: -1.372%; 13: -4.056%; 14: -3.72%; 17: -1.225%; 18: -1.708%; 19: -22.17%; 20: 14.06%; 21: -5.398%; 23: 23.45%; 24: 118.7%; 25: -3.226%
| Z-Score (Absolute) |  | 02: 0.38; 04: 0.94; 05: 2.87; 07: 0.67; 10: 0.03; 13: 0.58; 14: 0.51; 17: 0; 18: 0.1; 19: 4.31; 20: 3.15; 21: 0.86; 23: 5.08; 24: 24.7; 25: 0.41
| Analytical results for Te in OREAS 45f (Indicative Value 0.032 ppm) | [31 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.05 | < 0.05 | 0.03 | < 0.1 | 0.03 | 0.03 | 13 | < 0.1 | < 0.1 | < 0.02 | < 0.2 | < 0.05 | 0.056224 | < 0.05 | 0.12 | 0.02 | 0.05 | 0.02 | 0.14 | < 0.1 | < 0.05 | 0.02 | 0.07 | < 0.05 | < 0.01 | 0.02356 |  |  |  |  | 1
| 1 | 2 | < 0.05 | < 0.05 | 0.02 | < 0.1 | 0.02 | 0.02 | 15 | < 0.1 | < 0.1 | 0.03 | < 0.2 | < 0.05 | 0.047294 | < 0.05 | 0.11 | 0.04 | 0.07 | 0.01 | 0.11 | < 0.1 | < 0.05 | 0.03 | 0.07 | < 0.05 | 0.012 | 0.02641 |  |  |  |  | 9
| 1 | 3 | < 0.05 | < 0.05 | 0.02 | < 0.1 | 0.01 | 0.03 | 13 | < 0.1 | < 0.1 | < 0.02 | < 0.2 | < 0.05 | 0.062301 | < 0.05 | 0.11 | 0.05 | 0.06 | 0.02 | 0.14 | < 0.1 | 0.14 | 0.04 | 0.06 | < 0.05 | < 0.01 | 0.02816 |  |  |  |  | 16
| 1 | 4 | < 0.05 | < 0.05 | 0.04 | < 0.1 | 0.03 | 0.02 | 14 | < 0.1 | < 0.1 | < 0.02 | < 0.2 | < 0.05 | 0.053412 | < 0.05 | 0.11 | 0.03 | 0.06 | 0.02 | 0.12 | < 0.1 | < 0.05 | 0.04 | 0.07 | < 0.05 | 0.011 | 0.02498 |  |  |  |  | 0.03218
| 1 | 5 | < 0.05 | < 0.05 | 0.02 | < 0.1 | 0.02 | 0.02 | 13 | < 0.1 | < 0.1 | 0.02 | < 0.2 | < 0.05 | 0.045002 | < 0.05 | 0.12 | 0.04 | 0.05 | 0.02 | 0.13 | < 0.1 | < 0.05 | 0.03 | 0.05 | < 0.05 | 0.011 | 0.02759 |  |  |  |  | 29
| 1 | 6 | < 0.05 | < 0.05 | 0.02 | < 0.1 | 0.03 | 0.02 | 13 | < 0.1 | < 0.1 | 0.02 | < 0.2 | < 0.05 | 0.036019 | < 0.05 | 0.11 | 0.02 | 0.06 | 0.02 | 0.14 | < 0.1 | < 0.05 | 0.02 | 0.06 | < 0.05 | < 0.01 | 0.02804
| Mean |  | 05: 0.025; 10: 0.023333; 13: 0.023333; 14: 13.5; 19: 0.023333; 23: 0.050042; 25: 0.11333; 27: 0.033333; 28: 0.058333; 29: 0.018333; 30: 0.13; 33: 0.14; 34: 0.03; 35: 0.063333; 37: 0.011333; 38: 0.026457
| Median |  | 05: 0.02; 10: 0.025; 13: 0.02; 14: 13; 19: 0.02; 23: 0.050353; 25: 0.11; 27: 0.035; 28: 0.06; 29: 0.02; 30: 0.135; 33: 0.14; 34: 0.03; 35: 0.065; 37: 0.011; 38: 0.027
| Std Dev. |  | 05: 0.0083666; 10: 0.008165; 13: 0.005164; 14: 0.83666; 19: 0.0057735; 23: 0.0092679; 25: 0.005164; 27: 0.012111; 28: 0.0075277; 29: 0.0040825; 30: 0.012649; 34: 0.0089443; 35: 0.008165; 37: 0.00057735; 38: 0.0018583
| Rel.Std.Dev. |  | 05: 33.47%; 10: 34.99%; 13: 22.13%; 14: 6.197%; 19: 24.74%; 23: 18.52%; 25: 4.556%; 27: 36.33%; 28: 12.9%; 29: 22.27%; 30: 9.73%; 34: 29.81%; 35: 12.89%; 37: 5.094%; 38: 7.024%
| PDM3 |  | 05: -22.31%; 10: -27.49%; 13: -27.49%; 14: 41851%; 19: -27.49%; 23: 55.5%; 25: 252.2%; 27: 3.583%; 28: 81.27%; 29: -43.03%; 30: 304%; 33: 335%; 34: -6.776%; 35: 96.81%; 37: -64.78%; 38: -17.79%
| Z-Score (Absolute) |  | 0.2 | 0.2 | 0.2 | 1.37 | 0.31 | 0.31 | NA | 1.37 | 1.37 | 0.73 | 4.51 | 0.2 | 1.37 | 0.2 | 5.35 | 0.32 | 1.89 | 0.62 | 6.4 | NA | 1 | 0.11 | 2.21 | 0.2 | 1.26 | 0.11
Results from laboratories 17 & 33 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Th in OREAS 45f (Certified Value 7.67 ppm) | [31 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 8.8846 | 7.7 | 7.9 | 7.98 | 7.5068 | 7.7 | 7.82 | 8.8 | 6.88 | 7.2 | 8.6 | 5.8 | 7.44 | 7.1507 | 7.9 | 8.6 | 7.53 | 6.82 | 7.7 | 6.93 | 8.22 | 6.2 | 7 | 8.3 | 8.1267 | 4.164 | 7.2219 |  |  |  | 1
| 1 | 2 | 8.6981 | 7.6 | 8 | 8.24 | 7.0886 | 7.6 | 8.03 | 8.4 | 6.99 | 7.2 | 9 | 5.8 | 7.81 | 7.0928 | 7.91 | 8.8 | 7.78 | 6.85 | 7.3 | 6.86 | 7.5 | 6.3 | 6.88 | 8.3 | 7.9633 | 3.742 | 7.242 |  |  |  | #N/A
| 1 | 3 | 8.8807 | 7.5 | 8 | 7.74 | 7.2374 | 7.9 | 8.12 | 8.4 | 6.91 | 7.4 | 8.4 | 6.4 | 7.47 | 7.3782 | 7.91 | 8.6 | 8.29 | 7 | 7.1 | 6.76 | 8.05 | 5.8 | 7.54 | 8.2 | 8.1367 | 4.145 | 7.1099 |  |  |  | 16
| 1 | 4 | 8.7611 | 7.5 | 8.1 | 7.87 | 7.1187 | 7.8 | 7.97 | 8.7 | 7.12 | 7.2 | 8.9 | 5.8 | 7.28 | 7.5847 | 7.9 | 8.3 | 8.13 | 6.86 | 7.6 | 6.88 | 7.36 | 6 | 7.89 | 8.3 | 8.13 | 4.078 | 7.2648 |  |  |  | 7.6691
| 1 | 5 | 8.6608 | 7.5 | 8.1 | 7.64 | 7.5953 | 7.6 | 7.85 | 8.3 | 6.44 | 7.2 | 9 | 5.9 | 7.34 | 7.6353 | 7.93 | 8.6 | 8.28 | 6.84 | 7.2 | 6.9 | 8.24 | 6.2 | 7.65 | 8.2 | 8.1467 | 4.162 | 7.2126 |  |  |  | 181
| 1 | 6 | 8.6927 | 7.6 | 8 | 7.92 | 7.0974 | 7.4 | 7.88 | 8.6 | 6.2 | 7.2 | 8.7 | 6 | 7.47 | 7.2227 | 7.78 | 8.8 | 7.97 | 6.8 | 7.3 | 6.85 | 7.58 | 6.2 | 7.5 | 8.4 | 7.9467 | 4.042 | 7.1292
| Mean |  | 8.763 | 7.5667 | 8.0167 | 7.8983 | 7.274 | 7.6667 | 7.945 | 8.5333 | 6.7567 | 7.2333 | 8.7667 | 5.95 | 7.4683 | 7.3441 | 7.8883 | 8.6167 | 7.9967 | 6.8617 | 7.3667 | 6.8633 | 7.825 | 6.1167 | 7.41 | 8.2833 | 8.075 | 4.0555 | 7.1967
| Median |  | 8.7296 | 7.55 | 8 | 7.895 | 7.178 | 7.65 | 7.925 | 8.5 | 6.895 | 7.2 | 8.8 | 5.85 | 7.455 | 7.3005 | 7.905 | 8.6 | 8.05 | 6.845 | 7.3 | 6.87 | 7.815 | 6.2 | 7.52 | 8.3 | 8.1283 | 4.1115 | 7.2173
| Std Dev. |  | 0.098227 | 0.08165 | 0.075277 | 0.20808 | 0.22292 | 0.17512 | 0.11606 | 0.19664 | 0.35646 | 0.08165 | 0.24221 | 0.23452 | 0.18411 | 0.22765 | 0.054191 | 0.18348 | 0.29984 | 0.07111 | 0.23381 | 0.058195 | 0.39006 | 0.18348 | 0.39038 | 0.075277 | 0.093351 | 0.1613 | 0.062687
| Rel.Std.Dev. |  | 1.121% | 1.079% | 0.939% | 2.634% | 3.065% | 2.284% | 1.461% | 2.304% | 5.276% | 1.129% | 2.763% | 3.942% | 2.465% | 3.1% | 0.687% | 2.129% | 3.75% | 1.036% | 3.174% | 0.8479% | 4.985% | 3% | 5.268% | 0.9088% | 1.156% | 3.977% | 0.8711%
| PDM3 |  | 14.26% | -1.336% | 4.532% | 2.989% | -5.152% | -0.03191% | 3.597% | 11.27% | -11.9% | -5.682% | 14.31% | -22.42% | -2.618% | -4.238% | 2.858% | 12.36% | 4.271% | -10.53% | -3.944% | -10.51% | 2.033% | -20.24% | -3.379% | 8.009% | 5.292% | -47.12% | -6.159%
| Z-Score (Absolute) |  | 1.88 | 0 | 0.71 | 0.52 | 0.46 | 0.16 | 0.59 | 1.52 | 1.27 | 0.52 | 1.88 | 2.54 | 0.15 | 0.35 | 0.5 | 1.65 | 0.67 | 1.11 | 0.31 | 1.1 | 0.41 | 2.27 | 0.25 | 1.12 | 0.8 | 5.51 | 0.58
| Analytical results for Ti in OREAS 45f (Certified Value 0.097 wt.%) | [31 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 19 | 20 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | [5 empty] | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | [5 empty] | 3
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.1 | 0.12 | 0.16 | 0.12992 | 0.11 | 0.097 | 0.1 | 0.03 | 0.065 | 0.079 | 0.0666 | 0.18645 | 0.119 | 0.109 | 0.102 | 0.092 | 0.095 | 0.0935 | 0.08 | 0.092 | 0.164 | 0.08 | 0.1675 | 0.033972 | 0.085252 | [5 empty] | 1
| 1 | 2 | 0.11 | 0.13 | 0.166 | 0.11033 | 0.1 | 0.096 | 0.1 | 0.03 | 0.067 | 0.077 | 0.069 | 0.18818 | 0.116 | 0.112 | 0.101 | 0.096 | 0.095 | 0.092 | 0.06 | 0.092 | 0.136 | 0.08 | 0.15933 | 0.032036 | 0.07436 | [5 empty] | #N/A
| 1 | 3 | 0.115 | 0.14 | 0.162 | 0.12128 | 0.1 | 0.107 | 0.1 | 0.03 | 0.064 | 0.078 | 0.0672 | 0.19961 | 0.14 | 0.117 | 0.104 | 0.092 | 0.096 | 0.086 | 0.07 | 0.097 | 0.128 | 0.08 | 0.14527 | 0.034031 | 0.080652 | [5 empty] | 16
| 1 | 4 | 0.11 | 0.13 | 0.16 | 0.11295 | 0.11 | 0.1 | 0.1 | 0.03 | 0.067 | 0.078 | 0.0639 | 0.20347 | 0.127 | 0.113 | 0.103 | 0.089 | 0.092 | 0.0885 | 0.08 | 0.101 | 0.13 | 0.08 | 0.16177 | 0.032941 | 0.084103 | [5 empty] | 0.097419
| 1 | 5 | 0.115 | 0.13 | 0.162 | 0.12825 | 0.09 | 0.095 | 0.1 | 0.03 | 0.066 | 0.078 | 0.0626 | 0.20031 | 0.135 | 0.115 | 0.101 | 0.086 | 0.092 | 0.0955 | 0.08 | 0.091 | 0.123 | 0.08 | 0.1518 | 0.033528 | 0.086727 | [5 empty] | 182
| 1 | 6 | 0.115 | 0.13 | 0.163 | 0.12163 | 0.06 | 0.096 | 0.1 | 0.03 | 0.062 | 0.077 | 0.0639 | 0.19563 | 0.119 | 0.109 | 0.104 | 0.095 | 0.089 | 0.09 | 0.09 | 0.093 | 0.131 | 0.08 | 0.15783 | 0.034072 | 0.089742
| Mean |  | 0.11083 | 0.13 | 0.16217 | 0.12073 | 0.095 | 0.0985 | 0.1 | 0.03 | 0.065167 | 0.077833 | 0.065533 | 0.19561 | 0.126 | 0.1125 | 0.1025 | 0.091667 | 0.093167 | 0.090917 | 0.076667 | 0.094333 | 0.13533 | 0.08 | 0.15725 | 0.03343 | 0.083473
| Median |  | 0.1125 | 0.13 | 0.162 | 0.12145 | 0.1 | 0.0965 | 0.1 | 0.03 | 0.0655 | 0.078 | 0.06525 | 0.19762 | 0.123 | 0.1125 | 0.1025 | 0.092 | 0.0935 | 0.091 | 0.08 | 0.0925 | 0.1305 | 0.08 | 0.15858 | 0.03375 | 0.084678
| Std Dev. |  | 0.0058452 | 0.0063246 | 0.0022286 | 0.0078837 | 0.018708 | 0.0045056 | 1.5202e-17 | 0 | 0.0019408 | 0.00075277 | 0.0024443 | 0.0069153 | 0.009757 | 0.0032094 | 0.0013784 | 0.0037238 | 0.0026394 | 0.0034557 | 0.010328 | 0.0038816 | 0.014665 | 0 | 0.0077903 | 0.00080744 | 0.005375
| Rel.Std.Dev. |  | 5.274% | 4.865% | 1.374% | 6.53% | 19.69% | 4.574% | 1.52e-14% | 0% | 2.978% | 0.9672% | 3.73% | 3.535% | 7.744% | 2.853% | 1.345% | 4.062% | 2.833% | 3.801% | 13.47% | 4.115% | 10.84% | 0% | 4.954% | 2.415% | 6.439%
| PDM3 |  | 13.77% | 33.44% | 66.46% | 23.92% | -2.483% | 1.11% | 2.649% | -69.21% | -33.11% | -20.1% | -32.73% | 100.8% | 29.34% | 15.48% | 5.216% | -5.905% | -4.365% | -6.675% | -21.3% | -3.167% | 38.92% | -17.88% | 61.42% | -65.68% | -14.32%
| Z-Score (Absolute) |  | 0.61 | 1.35 | 2.59 | 0.99 | 0 | 0.13 | 0.19 | 2.5 | 1.15 | 0.66 | 1.14 | 3.88 | 1.19 | 0.67 | 0.29 | 0.13 | 0.07 | 0.16 | 0.71 | 0.03 | 1.55 | 0.58 | 2.4 | 2.37 | 0.44
| Analytical results for Tl in OREAS 45f (Certified Value 0.12 ppm) | [31 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.1 | 0.13 | 0.12 | 0.13093 | 0.14 | 0.103 | < 10 | < 0.1 | 0.1 | 0.12 | < 0.1 | 0.11 | 0.13672 | 0.12 | 0.12 | 0.111 | 0.105 | 0.12 | 0.099 | < 0.1 | 0.12 | 0.118 | < 0.01 | 0.16 | 0.044 | 0.11187 |  |  |  |  | 1
| 1 | 2 | 0.11 | 0.12 | 0.12 | 0.11612 | 0.14 | 0.101 | < 10 | < 0.1 | 0.1 | 0.12 | < 0.1 | 0.11 | 0.14801 | 0.117 | 0.12 | 0.117 | 0.104 | 0.1 | 0.096 | < 0.1 | 0.12 | 0.104 | < 0.01 | 0.15 | 0.038 | 0.10981 |  |  |  |  | 10
| 1 | 3 | 0.11 | 0.12 | 0.12 | 0.12683 | 0.14 | 0.102 | < 10 | < 0.1 | 0.1 | 0.11 | < 0.1 | 0.11 | 0.15885 | 0.145 | 0.12 | 0.12 | 0.108 | 0.1 | 0.097 | < 0.1 | 0.12 | 0.111 | < 0.01 | 0.15 | 0.046 | 0.11349 |  |  |  |  | 16
| 1 | 4 | 0.11 | 0.13 | 0.13 | 0.11933 | 0.15 | 0.102 | < 10 | < 0.1 | 0.1 | 0.11 | < 0.1 | 0.1 | 0.14043 | 0.129 | 0.12 | 0.12 | 0.104 | 0.11 | 0.097 | < 0.1 | 0.11 | 0.117 | < 0.01 | 0.15333 | 0.041 | 0.11226 |  |  |  |  | 0.11538
| 1 | 5 | 0.11 | 0.12 | 0.12 | 0.14875 | 0.14 | 0.101 | < 10 | < 0.1 | < 0.1 | 0.11 | < 0.1 | 0.11 | 0.15686 | 0.134 | 0.12 | 0.126 | 0.104 | 0.1 | 0.096 | < 0.1 | 0.11 | 0.11 | < 0.01 | 0.15667 | 0.043 | 0.11368 |  |  |  |  | 183
| 1 | 6 | 0.11 | 0.12 | 0.12 | 0.1424 | 0.11 | 0.098 | < 10 | < 0.1 | 0.1 | 0.11 | < 0.1 | 0.11 | 0.15429 | 0.12 | 0.12 | 0.117 | 0.104 | 0.11 | 0.098 | < 0.1 | 0.11 | 0.109 | < 0.01 | 0.15 | 0.042 | 0.11061
| Mean |  | 0.10833 | 0.12333 | 0.12167 | 0.13073 | 0.13667 | 0.10117 |  |  | 0.1 | 0.11333 |  | 0.10833 | 0.14919 | 0.1275 | 0.12 | 0.1185 | 0.10483 | 0.10667 | 0.097167 |  | 0.115 | 0.1115 |  | 0.15333 | 0.042333 | 0.11195
| Median |  | 0.11 | 0.12 | 0.12 | 0.12888 | 0.14 | 0.1015 |  |  | 0.1 | 0.11 |  | 0.11 | 0.15115 | 0.1245 | 0.12 | 0.1185 | 0.104 | 0.105 | 0.097 |  | 0.115 | 0.1105 |  | 0.15167 | 0.0425 | 0.11206
| Std Dev. |  | 0.0040825 | 0.005164 | 0.0040825 | 0.012801 | 0.013663 | 0.0017224 |  |  | 0 | 0.005164 |  | 0.0040825 | 0.009075 | 0.01071 | 0 | 0.0049295 | 0.0016021 | 0.008165 | 0.001169 |  | 0.0054772 | 0.005244 |  | 0.0042164 | 0.0027325 | 0.001539
| Rel.Std.Dev. |  | 3.768% | 4.187% | 3.355% | 9.792% | 9.997% | 1.703% |  |  | 0% | 4.556% |  | 3.768% | 6.083% | 8.4% | 0% | 4.16% | 1.528% | 7.655% | 1.203% |  | 4.763% | 4.703% |  | 2.75% | 6.455% | 1.375%
| PDM3 |  | -6.104% | 6.897% | 5.453% | 13.31% | 18.45% | -12.32% |  |  | -13.33% | -1.77% |  | -6.104% | 29.31% | 10.51% | 4.008% | 2.708% | -9.137% | -7.548% | -15.78% |  | -0.3254% | -3.359% |  | 32.9% | -63.31% | -2.966%
| Z-Score (Absolute) |  | 0.26 | 0.65 | 0.55 | 1.1 | 1.46 | 0.7 | 297.35 | 3.81 | NA | 0.04 | 3.81 | 0.26 | 2.22 | 0.9 | 0.45 | 0.36 | 0.48 | 0.36 | 0.94 | NA | 0.14 | 0.07 | 6.55 | 2.48 | 4.28 | 0.04
Results from laboratory 19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 45f (Certified Value 0.11 ppm) | [31 empty] | Certified
Tm | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 07: Lab; 10: Lab; 13: Lab; 14: Lab; 17: Lab |  |  | 1
| No | No | 02: 02; 04: 07; 05: 13; 07: 23; 10: 24; 13: 38; 14: 39; 17: 40 |  |  | ppm
|  |  | 02: AR*MS; 04: AR*OES/MS; 05: AR*OES/MS; 07: AR*MS; 10: AR*OES/MS; 13: AR*MS; 14: AR*MS; 17: AR*MS |  |  | 2
|  |  | 02: 0.5g; 04: 01g; 05: 25g; 07: 0.5g; 10: 01g; 13: 01g; 14: 0.5g; 17: 0.5g |  |  | 3
| 1 | 1 | 02: < 0.5; 04: 0.1; 05: 0.12; 07: 0.11; 10: 0.093227; 13: 0.11667; 14: 0.051; 17: 0.1004 |  |  | 1
| 1 | 2 | 02: < 0.5; 04: 0.11; 05: 0.11; 07: 0.11; 10: 0.098229; 13: 0.11; 14: < 0.05; 17: 0.10424 |  |  | #N/A
| 1 | 3 | 02: < 0.5; 04: 0.1; 05: 0.12; 07: 0.11; 10: 0.10069; 13: 0.11667; 14: 0.051; 17: 0.09896 |  |  | 16
| 1 | 4 | 02: < 0.5; 04: 0.1; 05: 0.12; 07: 0.1; 10: 0.10017; 13: 0.11667; 14: < 0.05; 17: 0.10016 |  |  | 0.10731
| 1 | 5 | 02: < 0.5; 04: 0.1; 05: 0.12; 07: 0.11; 10: 0.10072; 13: 0.11667; 14: 0.051; 17: 0.10202 |  |  | 184
| 1 | 6 | 02: < 0.5; 04: 0.11; 05: 0.11; 07: 0.11; 10: 0.094645; 13: 0.11667; 14: < 0.05; 17: 0.09981
| Mean |  | 04: 0.10333; 05: 0.11667; 07: 0.10833; 10: 0.097946; 13: 0.11556; 14: 0.051; 17: 0.10093
| Median |  | 04: 0.1; 05: 0.12; 07: 0.11; 10: 0.099201; 13: 0.11667; 14: 0.051; 17: 0.10028
| Std Dev. |  | 04: 0.005164; 05: 0.005164; 07: 0.0040825; 10: 0.0032677; 13: 0.0027217; 14: 0; 17: 0.0019054
| Rel.Std.Dev. |  | 04: 4.997%; 05: 4.426%; 07: 3.768%; 10: 3.336%; 13: 2.355%; 14: 0%; 17: 1.888%
| PDM3 |  | 04: -3.708%; 05: 8.716%; 07: 0.9508%; 10: -8.729%; 13: 7.681%; 14: -52.48%; 17: -5.946%
| Z-Score (Absolute) |  | 02: 11.04; 04: 0.19; 05: 0.83; 07: 0.19; 10: 0.6; 13: 0.74; 14: 5.19; 17: 0.38
| Analytical results for U in OREAS 45f (Certified Value 1.09 ppm) | [31 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.4g | 40g | 15g | 0.5g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 1.2149 | 1.1 | 1.12 | 1.1 | 0.58269 | 1.1 | 1.09 | 0.94 | 1 | 1.2 | 0.9 | 1.03 | 1.0395 | 1.06 | 1.29 | 1.05 | 1.015 | 1.07 | 1.025 | 1.26 | 1.13 | 0.973 | 1.18 | 1.1733 | 0.512 | 1.0322 |  |  |  |  | 1
| 1 | 2 | 1.2043 | 1.1 | 1.13 | 1.14 | 0.51707 | 1.09 | 1.08 | 0.96 | 1 | 1.3 | 0.9 | 1.07 | 1.0565 | 1.07 | 1.23 | 1.175 | 1.02 | 1.02 | 1.03 | 1.08 | 1.15 | 0.944 | 1.2 | 1.13 | 0.457 | 1.0365 |  |  |  |  | #N/A
| 1 | 3 | 1.2187 | 1.06 | 1.13 | 1.08 | 0.56154 | 1.12 | 1.11 | 0.95 | 1 | 1.2 | 1 | 1.04 | 1.0779 | 1.1 | 1.2 | 1.16 | 1.04 | 0.98 | 1.01 | 1.18 | 1.08 | 1.03 | 1.19 | 1.1433 | 0.505 | 1.0334 |  |  |  |  | 16
| 1 | 4 | 1.2087 | 1.08 | 1.16 | 1.1 | 0.52565 | 1.11 | 1.065 | 0.97 | 1 | 1.3 | 0.9 | 1.02 | 1.1259 | 1.1 | 1.2 | 1.12 | 1.02 | 1.03 | 1.025 | 1.09 | 1.09 | 1.075 | 1.18 | 1.1333 | 0.491 | 1.0385 |  |  |  |  | 1.0877
| 1 | 5 | 1.2015 | 1.06 | 1.16 | 1.07 | 0.59264 | 1.1 | 1.085 | 0.88 | 1 | 1.3 | 0.9 | 1.03 | 1.132 | 1.08 | 1.23 | 1.155 | 1.01 | 0.98 | 1.03 | 1.21 | 1.13 | 1.045 | 1.18 | 1.1467 | 0.493 | 1.0199 |  |  |  |  | 185
| 1 | 6 | 1.2076 | 1.06 | 1.14 | 1.11 | 0.56604 | 1.03 | 1.055 | 0.85 | 1 | 1.3 | 1 | 1.04 | 1.0515 | 1.05 | 1.25 | 1.1 | 1.015 | 1.02 | 1.04 | 1.11 | 1.12 | 1.04 | 1.21 | 1.1233 | 0.482 | 1.0192
| Mean |  | 1.2093 | 1.0767 | 1.14 | 1.1 | 0.5576 | 1.0917 | 1.0808 | 0.925 | 1 | 1.2667 | 0.93333 | 1.0383 | 1.0805 | 1.0767 | 1.2333 | 1.1267 | 1.02 | 1.0167 | 1.0267 | 1.155 | 1.1167 | 1.0178 | 1.19 | 1.1417 | 0.49 | 1.03
| Median |  | 1.2082 | 1.07 | 1.135 | 1.1 | 0.56379 | 1.1 | 1.0825 | 0.945 | 1 | 1.3 | 0.9 | 1.035 | 1.0672 | 1.075 | 1.23 | 1.1375 | 1.0175 | 1.02 | 1.0275 | 1.145 | 1.125 | 1.035 | 1.185 | 1.1383 | 0.492 | 1.0328
| Std Dev. |  | 0.0064623 | 0.019664 | 0.016733 | 0.024495 | 0.030355 | 0.031885 | 0.019343 | 0.048477 | 0 | 0.05164 | 0.05164 | 0.017224 | 0.03955 | 0.020656 | 0.033862 | 0.046655 | 0.010488 | 0.033862 | 0.0098319 | 0.07287 | 0.026583 | 0.049207 | 0.012649 | 0.017733 | 0.019349 | 0.0083667
| Rel.Std.Dev. |  | 0.5344% | 1.826% | 1.468% | 2.227% | 5.444% | 2.921% | 1.79% | 5.241% | 0% | 4.077% | 5.533% | 1.659% | 3.66% | 1.919% | 2.746% | 4.141% | 1.028% | 3.331% | 0.9577% | 6.309% | 2.381% | 4.835% | 1.063% | 1.553% | 3.949% | 0.8123%
| PDM3 |  | 11.17% | -1.018% | 4.804% | 1.127% | -48.74% | 0.3609% | -0.635% | -14.96% | -8.066% | 16.45% | -14.2% | -4.542% | -0.6621% | -1.018% | 13.38% | 3.579% | -6.228% | -6.534% | -5.615% | 6.183% | 2.659% | -6.427% | 9.401% | 4.958% | -54.95% | -5.312%
| Z-Score (Absolute) |  | 1.44 | 0.02 | 0.68 | 0.24 | 5.75 | 0.14 | 0.02 | 1.7 | 0.87 | 2.08 | 1.6 | 0.44 | 0.02 | 0.02 | 1.71 | 0.53 | 0.65 | 0.68 | 0.57 | 0.84 | 0.42 | 0.67 | 1.23 | 0.7 | 6.5 | 0.54
| Analytical results for V in OREAS 45f (Certified Value 217 ppm) | [31 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 15g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 216.1582 | 218 | 215 | 217 | 221.7608 | 223 | 211 | 240 | 196 | 201 | 196 | 220 | 209 | 214 | 235.2156 | 217 | 217 | 218 | 217 | 212 | 214 | 203 | 219 | 207 | 185 | 228 | 233.6667 | 158.739 | 237.3869 |  | 1
| 1 | 2 | 215.3303 | 213 | 218 | 216 | 235.4037 | 219 | 210 | 226 | 197 | 201 | 192 | 221 | 214 | 219 | 219.2577 | 215 | 220 | 225 | 218 | 223 | 215 | 199 | 220 | 201 | 228 | 229 | 231 | 152.725 | 229.8552 |  | #N/A
| 1 | 3 | 215.8742 | 209 | 225 | 214 | 211.3995 | 221 | 212 | 233 | 193 | 201 | 204 | 223 | 213 | 213 | 224.4322 | 225 | 217 | 219 | 217 | 223 | 216 | 197 | 210 | 209 | 240 | 231 | 232 | 160.77 | 237.7132 |  | 16
| 1 | 4 | 214.3409 | 203 | 228 | 215 | 226.7945 | 215 | 211 | 233 | 200 | 206 | 202 | 221 | 214 | 210 | 223.2717 | 216 | 222 | 219 | 218 | 218 | 216 | 199 | 216 | 209 | 225 | 224 | 234 | 157.67 | 230.3654 |  | 217.0678
| 1 | 5 | 219.1902 | 210 | 222 | 218 | 224.6532 | 222 | 208 | 229 | 197 | 204 | 200 | 223 | 214 | 205 | 232.2104 | 218 | 228 | 217 | 216 | 215 | 216 | 202 | 217 | 206 | 222 | 230 | 231.6667 | 160.132 | 232.6251 |  | 186
| 1 | 6 | 216.2608 | 212 | 221 | 215 | 209.7834 | 215 | 208 | 235 | 198 | 202 | 196 | 217 | 217 | 213 | 223.4635 | 217 | 236 | 212 | 216 | 226 | 212 | 205 | 218 | 209 | 229 | 227 | 232.3333 | 158.477 | 233.5689
| Mean |  | 216.1924 | 210.8333 | 221.5 | 215.8333 | 221.6325 | 219.1667 | 210 | 232.6667 | 196.8333 | 202.5 | 198.3333 | 220.8333 | 213.5 | 212.3333 | 226.3085 | 218 | 223.3333 | 218.3333 | 217 | 219.5 | 214.8333 | 200.8333 | 216.6667 | 206.8333 | 221.5 | 228.1667 | 232.4444 | 158.0855 | 233.5858
| Median |  | 216.0162 | 211 | 221.5 | 215.5 | 223.207 | 220 | 210.5 | 233 | 197 | 201.5 | 198 | 221 | 214 | 213 | 223.9478 | 217 | 221 | 218.5 | 217 | 220.5 | 215.5 | 200.5 | 217.5 | 208 | 226.5 | 228.5 | 232.1667 | 158.608 | 233.097
| Std Dev. |  | 1.6288 | 4.9565 | 4.6797 | 1.472 | 9.701 | 3.4881 | 1.6733 | 4.8442 | 2.3166 | 2.0736 | 4.4572 | 2.2286 | 2.5884 | 4.6332 | 6.0782 | 3.5777 | 7.4207 | 4.1793 | 0.89443 | 5.3944 | 1.6021 | 2.9944 | 3.559 | 3.1252 | 18.8971 | 2.4833 | 1.1675 | 2.8595 | 3.3674
| Rel.Std.Dev. |  | 0.7534% | 2.351% | 2.113% | 0.682% | 4.377% | 1.592% | 0.7968% | 2.082% | 1.177% | 1.024% | 2.247% | 1.009% | 1.212% | 2.182% | 2.686% | 1.641% | 3.323% | 1.914% | 0.4122% | 2.458% | 0.7457% | 1.491% | 1.643% | 1.511% | 8.531% | 1.088% | 0.5023% | 1.809% | 1.442%
| PDM3 |  | -0.4032% | -2.872% | 2.042% | -0.5687% | 2.103% | 0.9669% | -3.256% | 7.186% | -9.322% | -6.711% | -8.631% | 1.735% | -1.644% | -2.181% | 4.257% | 0.4295% | 2.886% | 0.583% | -0.03121% | 1.121% | -1.029% | -7.479% | -0.1848% | -4.715% | 2.042% | 5.113% | 7.084% | -27.17% | 7.61%
| Z-Score (Absolute) |  | 0.12 | 0.89 | 0.65 | 0.17 | 0.67 | 0.31 | 1.01 | 2.26 | 2.91 | 2.1 | 2.7 | 0.55 | 0.51 | 0.67 | 1.35 | 0.14 | 0.92 | 0.19 | 0 | 0.36 | 0.31 | 2.34 | 0.05 | 1.47 | 0.65 | 1.61 | 2.23 | 8.51 | 2.4
| Analytical results for W in OREAS 45f (Indicative Value 0.006 ppm) | [31 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.1 | < 0.1 | < 0.05 | 0.12347 | < 0.05 | 0.006 | < 5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 1.9 | 0.003 | 0.003 | < 0.05 | 0.009 | < 0.1 | < 0.1 | 0.007 | 0.05 | < 10 | < 0.05 | < 0.1 | 0.00631 |  |  |  |  | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.05 | 0.005 | < 5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 1.7 | 0.005 | 0.007 | < 0.05 | 0.01 | < 0.1 | < 0.1 | 0.005 | 0.1 | < 10 | < 0.05 | < 0.1 | 0.00826 |  |  |  |  | 11
| 1 | 3 | < 0.1 | < 0.1 | < 0.05 | 0.19657 | < 0.05 | 0.019 | < 5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 1.8 | 0.024 | 0.007 | < 0.05 | 0.011 | < 0.1 | < 0.1 | 0.006 | 0.07 | < 10 | < 0.05 | < 0.1 | 0.00825 |  |  |  |  | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | < 0.05 | 0.004 | < 5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 1.7 | 0.002 | 0.005 | < 0.05 | 0.019 | < 0.1 | < 0.1 | 0.008 | 0.07 | < 10 | < 0.05 | < 0.1 | 0.00825 |  |  |  |  | 0.0063381
| 1 | 5 | < 0.1 | < 0.1 | < 0.05 | 0.19417 | < 0.05 | 0.005 | < 5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 1.9 | 0.002 | 0.006 | < 0.05 | 0.011 | < 0.1 | < 0.1 | 0.006 | 0.09 | < 10 | < 0.05 | < 0.1 | 0.00567 |  |  |  |  | 30
| 1 | 6 | < 0.1 | < 0.1 | < 0.05 | 0.18868 | < 0.05 | 0.008 | < 5 | < 0.1 | < 0.5 | < 0.1 | < 0.1 | < 0.05 | < 0.1 | 1.7 | 0.002 | 0.005 | < 0.05 | 0.01 | < 0.1 | < 0.1 | 0.009 | 0.08 | < 10 | < 0.05 | < 0.1 | 0.00583
| Mean |  | 10: 0.17573; 14: 0.0078333; 27: 1.7833; 28: 0.0063333; 29: 0.0055; 31: 0.011667; 35: 0.0068333; 36: 0.076667; 40: 0.007095
| Median |  | 10: 0.19143; 14: 0.0055; 27: 1.75; 28: 0.0025; 29: 0.0055; 31: 0.0105; 35: 0.0065; 36: 0.075; 40: 0.00728
| Std Dev. |  | 10: 0.034992; 14: 0.0056362; 27: 0.098319; 28: 0.0087331; 29: 0.0015166; 31: 0.0036697; 35: 0.001472; 36: 0.017512; 40: 0.0012863
| Rel.Std.Dev. |  | 10: 19.91%; 14: 71.95%; 27: 5.513%; 28: 137.9%; 29: 27.57%; 31: 31.45%; 35: 21.54%; 36: 22.84%; 40: 18.13%
| PDM3 |  | 10: 2673%; 14: 23.59%; 27: 28037%; 28: -0.07451%; 29: -13.22%; 31: 84.07%; 35: 7.814%; 36: 1110%; 40: 11.94%
| Z-Score (Absolute) |  | 0 | 0 | 0.67 | 2.26 | 0.67 | 1.14 | 66.08 | 0 | 5.39 | 0 | 0 | 0.67 | 0 | 46.75 | 1.18 | 1.2 | 0.67 | 1.03 | 0 | 0 | 1.16 | 0.72 | 133.51 | 0.67 | 0 | 1.16
| Analytical results for Y in OREAS 45f (Certified Value 6.74 ppm) | [31 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 0.15g | 0.5g | 01g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 7.7389 | 6.99 | 6.85 | 6.73 | 2175.649 | 7.15 | 6.65 | 7 | 6 | 6.19 | 6.8 | 6 | 6.66 | 6.3254 | 6.63 | 5.61 | 6.77 | 6.61 | 6.54 | 7.33 | 6.54 | 6.98 | 6.49 | 7.17 | 7.4 | 7.87 | 4.453 | 6.4002 |  |  | 1
| 1 | 2 | 7.5945 | 6.92 | 6.75 | 6.7 | 2041.653 | 6.71 | 6.64 | 7 | 6 | 6.13 | 6.7 | 7 | 6.91 | 6.221 | 6.65 | 5.98 | 6.82 | 6.75 | 6.26 | 7.06 | 6.67 | 7.08 | 6.52 | 7.29 | 7.2 | 7.7567 | 4.472 | 6.6321 |  |  | #N/A
| 1 | 3 | 7.7695 | 6.75 | 6.87 | 6.64 | 2087.5347 | 6.98 | 6.73 | 7 | 6 | 6.16 | 6.8 | 6 | 6.64 | 6.3546 | 6.69 | 5.81 | 7.1 | 6.7 | 6.17 | 6.89 | 6.42 | 6.75 | 7.12 | 7.28 | 7.6 | 7.5633 | 4.496 | 6.5332 |  |  | 16
| 1 | 4 | 7.6897 | 6.77 | 6.93 | 6.79 | 2064.3591 | 7.01 | 6.63 | 7 | 6 | 6.36 | 6.7 | 7 | 6.51 | 6.3908 | 6.7 | 5.89 | 6.87 | 6.77 | 6.43 | 6.84 | 7.13 | 6.76 | 7.28 | 7.28 | 7.3 | 7.7967 | 4.52 | 6.6895 |  |  | 6.7424
| 1 | 5 | 7.7006 | 6.76 | 6.88 | 6.14 | 2107.6922 | 6.82 | 6.45 | 7 | 6 | 5.98 | 6.6 | 6 | 6.55 | 6.4187 | 6.57 | 5.92 | 6.97 | 6.69 | 6.3 | 6.98 | 6.55 | 6.92 | 6.98 | 7.43 | 7.5 | 7.7033 | 4.509 | 6.5977 |  |  | 187
| 1 | 6 | 7.7079 | 6.75 | 6.79 | 6.84 | 1903.2813 | 6.55 | 6.49 | 7 | 6 | 5.75 | 6.7 | 7 | 6.68 | 6.213 | 6.62 | 6.14 | 6.64 | 6.61 | 6.44 | 6.97 | 6.9 | 6.92 | 6.99 | 7.11 | 7.4 | 7.6767 | 4.45 | 6.3875
| Mean |  | 7.7002 | 6.8233 | 6.845 | 6.64 | 2063.3615 | 6.87 | 6.5983 | 7 | 6 | 6.095 | 6.7167 | 6.5 | 6.6583 | 6.3206 | 6.6433 | 5.8917 | 6.8617 | 6.6883 | 6.3567 | 7.0117 | 6.7017 | 6.9017 | 6.8967 | 7.26 | 7.4 | 7.7278 | 4.4833 | 6.54
| Median |  | 7.7043 | 6.765 | 6.86 | 6.715 | 2075.9469 | 6.9 | 6.635 | 7 | 6 | 6.145 | 6.7 | 6.5 | 6.65 | 6.34 | 6.64 | 5.905 | 6.845 | 6.695 | 6.365 | 6.975 | 6.61 | 6.92 | 6.985 | 7.28 | 7.4 | 7.73 | 4.484 | 6.5655
| Std Dev. |  | 0.059412 | 0.10463 | 0.065038 | 0.25464 | 90.8361 | 0.21936 | 0.10629 | 0 | 0 | 0.2083 | 0.075277 | 0.54772 | 0.13992 | 0.086295 | 0.048028 | 0.1768 | 0.15993 | 0.067651 | 0.13663 | 0.1736 | 0.26544 | 0.12781 | 0.32241 | 0.11063 | 0.14142 | 0.10589 | 0.029405 | 0.12408
| Rel.Std.Dev. |  | 0.7716% | 1.533% | 0.9502% | 3.835% | 4.402% | 3.193% | 1.611% | 0% | 0% | 3.418% | 1.121% | 8.427% | 2.101% | 1.365% | 0.7229% | 3.001% | 2.331% | 1.011% | 2.149% | 2.476% | 3.961% | 1.852% | 4.675% | 1.524% | 1.911% | 1.37% | 0.6559% | 1.897%
| PDM3 |  | 14.21% | 1.201% | 1.522% | -1.518% | 30503% | 1.893% | -2.136% | 3.821% | -11.01% | -9.601% | -0.3811% | -3.595% | -1.246% | -6.256% | -1.469% | -12.62% | 1.769% | -0.8013% | -5.72% | 3.994% | -0.6036% | 2.363% | 2.289% | 7.677% | 9.754% | 14.62% | -33.5% | -3.001%
| Z-Score (Absolute) |  | 3.33 | 0.38 | 0.46 | 0.23 | 6905.32 | 0.54 | 0.37 | 0.98 | 2.38 | 2.06 | 0.03 | 0.7 | 0.17 | 1.3 | 0.22 | 2.74 | 0.51 | 0.07 | 1.18 | 1.02 | 0.03 | 0.65 | 0.63 | 1.85 | 2.32 | 3.42 | 7.47 | 0.57
| Analytical results for Yb in OREAS 45f (Certified Value 0.69 ppm) | [31 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 02 | 04 | 05 | 07 | 13 | 15 | 18 | 23 | 24 | 25 | 27 | 33 | 34 | 38 | 39 | 40 | [14 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [14 empty] | 2
|  |  | 0.5g | 0.25g | 0.25g | 01g | 25g | 0.25g | 0.4g | 0.5g | 01g | 0.25g | 0.25g | 0.4g | 0.25g | 01g | 0.5g | 0.5g | [14 empty] | 3
| 1 | 1 | 0.7353 | 0.7 | 0.7 | 0.66 | 0.7 | 1.7 | 0.7 | 0.66 | 0.60303 | 0.686 | 0.7 | 0.8 | 0.7 | 0.78 | 0.301 | 0.63534 | [14 empty] | 1
| 1 | 2 | 0.7387 | 0.7 | 0.7 | 0.66 | 0.68 | 1.6 | 0.7 | 0.68 | 0.63058 | 0.67 | 0.7 | 0.7 | 0.7 | 0.74333 | 0.299 | 0.6499 | [14 empty] | #N/A
| 1 | 3 | 0.7508 | 0.7 | 0.7 | 0.64 | 0.72 | 1.7 | 0.7 | 0.66 | 0.63505 | 0.656 | 0.7 | 0.7 | 0.6 | 0.76667 | 0.32 | 0.63697 | [14 empty] | 16
| 1 | 4 | 0.747 | 0.7 | 0.7 | 0.66 | 0.65 | 1.7 | 0.7 | 0.65 | 0.67118 | 0.659 | 0.7 | 0.8 | 0.7 | 0.75667 | 0.299 | 0.63788 | [14 empty] | 0.68935
| 1 | 5 | 0.7424 | 0.7 | 0.7 | 0.62 | 0.67 | 1.7 | 0.7 | 0.65 | 0.65465 | 0.664 | 0.7 | 0.7 | 0.7 | 0.77 | 0.311 | 0.65913 | [14 empty] | 188
| 1 | 6 | 0.7579 | 0.6 | 0.7 | 0.64 | 0.65 | 1.7 | 0.7 | 0.67 | 0.61324 | 0.661 | 0.7 | 0.8 | 0.7 | 0.75333 | 0.308 | 0.64459
| Mean |  | 0.74535 | 0.68333 | 0.7 | 0.64667 | 0.67833 | 1.6833 | 0.7 | 0.66167 | 0.63462 | 0.666 | 0.7 | 0.75 | 0.68333 | 0.76167 | 0.30633 | 0.64397
| Median |  | 0.7447 | 0.7 | 0.7 | 0.65 | 0.675 | 1.7 | 0.7 | 0.66 | 0.63282 | 0.6625 | 0.7 | 0.75 | 0.7 | 0.76167 | 0.3045 | 0.64123
| Std Dev. |  | 0.0082928 | 0.040825 | 1.2162e-16 | 0.01633 | 0.027869 | 0.040825 | 1.2162e-16 | 0.01169 | 0.025335 | 0.0109 | 1.2162e-16 | 0.054772 | 0.040825 | 0.013123 | 0.0083347 | 0.0092297
| Rel.Std.Dev. |  | 1.113% | 5.974% | 1.737e-14% | 2.525% | 4.108% | 2.425% | 1.737e-14% | 1.767% | 3.992% | 1.637% | 1.737e-14% | 7.303% | 5.974% | 1.723% | 2.721% | 1.433%
| PDM3 |  | 8.123% | -0.8732% | 1.545% | -6.192% | -1.599% | 144.2% | 1.545% | -4.016% | -7.939% | -3.388% | 1.545% | 8.798% | -0.8732% | 10.49% | -55.56% | -6.584%
| Z-Score (Absolute) |  | 1.43 | 0 | 0.39 | 0.85 | 0.12 | 23.12 | 0.39 | 0.5 | 1.13 | 0.4 | 0.39 | 1.54 | 0 | 1.81 | 8.72 | 0.91
| Analytical results for Zn in OREAS 45f (Certified Value 22.2 ppm) | [31 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 04 | 05 | 07 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.4g | 40g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g |  | 1
| 1 | 1 | 25.183 | 18.2 | 18 | 41 | 48.4968 | 30 | 21.3 | 30 | 66 | 15 | 20 | 21 | 16 | 20.6 | 33.6737 | 22.2 | 24.0803 | 23.3 | 20.3 | 24 | 21.3 | 24 | 24 | 21 | 24.3 | 24.4 | 27.3333 | 7.21 | 22.3039 |  | 1
| 1 | 2 | 25.29 | 17.3 | 18 | 41 | 47.6034 | 28 | 21.2 | 31 | 64 | 15 | 20 | 21 | 15 | 21.5 | 33.8485 | 20.8 | 22.2174 | 22.2 | 20.2 | 24 | 20.7 | 22 | 27 | 20.8 | 20 | 26.7 | 27 | 6.9 | 21.6461 |  | 1
| 1 | 3 | 24.688 | 17.4 | 19 | 40 | 52.2957 | 28 | 21.7 | 29 | 62 | 15 | 20 | 21 | 16 | 20.7 | 35.8057 | 22.7 | 22.2086 | 23.7 | 21 | 24 | 20.3 | 23 | 25 | 23.1 | 21 | 30.3 | 26.6667 | 7.07 | 22.7757 |  | 16
| 1 | 4 | 25.074 | 17.1 | 21 | 41 | 49.9971 | 29 | 21.2 | 28 | 69 | 15 | 20 | 20 | 17 | 19.9 | 36.1277 | 23.6 | 22.2631 | 23.2 | 20.2 | 23 | 21 | 25 | 26 | 23.1 | 20.6 | 26.2 | 27.6667 | 6.954 | 21.9722 |  | 22.2014
| 1 | 5 | 25.461 | 17.7 | 19 | 42 | 50.7508 | 28 | 20.9 | 32 | 63 | 15 | 20 | 21 | 16 | 20.3 | 36.5409 | 23.9 | 22.2649 | 22.9 | 20.4 | 23 | 20.3 | 20 | 26 | 22 | 20.1 | 27.3 | 26.6667 | 7.146 | 23.4393 |  | 189
| 1 | 6 | 24.746 | 17.2 | 19 | 42 | 48.4225 | 22 | 20.4 | 30 | 66 | 14 | 20 | 21 | 18 | 20.1 | 34.6301 | 21.4 | 22.1244 | 22.8 | 20 | 23 | 20 | 24 | 27 | 22.8 | 21 | 26.1 | 26.6667 | 7.17 | 22.3227
| Mean |  | 25.0737 | 17.4833 | 19 | 41.1667 | 49.5944 | 27.5 | 21.1167 | 30 | 65 | 14.8333 | 20 | 20.8333 | 16.3333 | 20.5167 | 35.1044 | 22.4333 | 22.5265 | 23.0167 | 20.35 | 23.5 | 20.6 | 23 | 25.8333 | 22.1333 | 21.1667 | 26.8333 | 27 | 7.075 | 22.41
| Median |  | 25.1285 | 17.35 | 19 | 41 | 49.247 | 28 | 21.2 | 30 | 65 | 15 | 20 | 21 | 16 | 20.45 | 35.2179 | 22.45 | 22.2403 | 23.05 | 20.25 | 23.5 | 20.5 | 23.5 | 26 | 22.4 | 20.8 | 26.45 | 26.8333 | 7.108 | 22.3133
| Std Dev. |  | 0.30492 | 0.40702 | 1.0954 | 0.75277 | 1.7515 | 2.8107 | 0.43551 | 1.4142 | 2.5298 | 0.40825 | 0 | 0.40825 | 1.0328 | 0.56716 | 1.2208 | 1.2144 | 0.76294 | 0.51153 | 0.34496 | 0.54772 | 0.4899 | 1.7889 | 1.169 | 1.0386 | 1.5933 | 1.9552 | 0.42164 | 0.12456 | 0.63043
| Rel.Std.Dev. |  | 1.216% | 2.328% | 5.766% | 1.829% | 3.532% | 10.22% | 2.062% | 4.714% | 3.892% | 2.752% | 0% | 1.96% | 6.323% | 2.764% | 3.478% | 5.413% | 3.387% | 2.222% | 1.695% | 2.331% | 2.378% | 7.778% | 4.525% | 4.692% | 7.527% | 7.286% | 1.562% | 1.761% | 2.813%
| PDM3 |  | 12.94% | -21.25% | -14.42% | 85.42% | 123.4% | 23.87% | -4.886% | 35.13% | 192.8% | -33.19% | -9.915% | -6.162% | -26.43% | -7.588% | 58.12% | 1.045% | 1.464% | 3.672% | -8.339% | 5.849% | -7.213% | 3.597% | 16.36% | -0.3064% | -4.661% | 20.86% | 21.61% | -68.13% | 0.9398%
| Z-Score (Absolute) |  | 0.67 | 1.26 | 0.88 | 4.78 | 6.94 | 1.29 | 0.34 | 1.93 | 10.87 | 1.94 | 0.62 | 0.41 | 1.56 | 0.49 | 3.24 | 0 | 0.02 | 0.15 | 0.53 | 0.27 | 0.47 | 0.14 | 0.87 | 0.08 | 0.32 | 1.12 | 1.17 | 3.92 | 0.01
| Analytical results for Zr in OREAS 45f (Certified Value 30 ppm) | [31 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 04 | 05 | 07 | 10 | 13 | 14 | 17 | 18 | 19 | 21 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 36 | 37 | 38 | 39 | 40 | [5 empty] | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | [5 empty] | 1
|  |  | 0.25g | 0.25g | 01g | 01g | 25g | 0.5g | 0.5g | 0.4g | 40g | 0.15g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.25g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | [5 empty] | 1
| 1 | 1 | 32.6 | 29.1 | 30.4 | 76.6828 | 21.6 | 25.6 | 34 | 29.7 | 19 | 28 | 23.26 | 36.1242 | 31 | 27.3 | 28.2 | 29.1 | 29.2 | 32.5 | 34.8 | 29.9 | 46.7 | 36.2 | 41.1333 | 25.548 | 29.0984 | [5 empty] | 1
| 1 | 2 | 30.3 | 29.4 | 31.9 | 69.2717 | 19.6 | 24.2 | 33 | 29.3 | 18.5 | 28 | 24.89 | 35.8112 | 32.6 | 28.1 | 28.6 | 27.9 | 28.1 | 30.3 | 32.8 | 29.8 | 49.2 | 35.8 | 37.9667 | 24.967 | 30.4304 | [5 empty] | #N/A
| 1 | 3 | 30.4 | 30 | 30.1 | 52.8572 | 22.2 | 27.3 | 33 | 28.8 | 20.5 | 29 | 23.84 | 36.9107 | 34.8 | 27.6 | 29 | 25.7 | 28.7 | 28.3 | 33.5 | 31.7 | 47.7 | 36.2 | 32.5667 | 26.146 | 30.4903 | [5 empty] | 16
| 1 | 4 | 31.1 | 30.2 | 31.3 | 52.4743 | 18.9 | 25.6 | 36 | 30 | 19 | 29 | 22.27 | 37.8066 | 32.8 | 26.6 | 28.5 | 27.7 | 29.9 | 31.3 | 34.2 | 32.8 | 47.7 | 35.5 | 38.7 | 25.372 | 30.3455 | [5 empty] | 30.0433
| 1 | 5 | 30.3 | 29 | 29.1 | 68.2892 | 21 | 24.1 | 34 | 29.2 | 19.5 | 29 | 22.41 | 37.9938 | 33.7 | 26.3 | 28 | 26.2 | 29.5 | 32.3 | 33.7 | 30.6 | 49.1 | 36.1 | 37.7667 | 26.277 | 29.924 | [5 empty] | 190
| 1 | 6 | 30.9 | 28.7 | 31.5 | 86.5614 | 24.4 | 24.2 | 34 | 28.4 | 19.5 | 29 | 23.2 | 36.9244 | 32.3 | 25.9 | 29.5 | 28.6 | 28.4 | 32.7 | 34.5 | 32.3 | 48.2 | 36.3 | 38.1 | 26.454 | 28.5251
| Mean |  | 30.9333 | 29.4 | 30.7167 | 67.6894 | 21.2833 | 25.1667 | 34 | 29.2333 | 19.3333 | 28.6667 | 23.3117 | 36.9285 | 32.8667 | 26.9667 | 28.6333 | 27.5333 | 28.9667 | 31.2333 | 33.9167 | 31.1833 | 48.1 | 36.0167 | 37.7056 | 25.794 | 29.8023
| Median |  | 30.65 | 29.25 | 30.85 | 68.7805 | 21.3 | 24.9 | 34 | 29.25 | 19.25 | 29 | 23.23 | 36.9175 | 32.7 | 26.95 | 28.55 | 27.8 | 28.95 | 31.8 | 33.95 | 31.15 | 47.95 | 36.15 | 38.0333 | 25.847 | 30.1348
| Std Dev. |  | 0.88242 | 0.58992 | 1.0439 | 13.3521 | 1.9621 | 1.2596 | 1.0954 | 0.58195 | 0.68313 | 0.5164 | 0.96783 | 0.87195 | 1.2894 | 0.83825 | 0.5465 | 1.3337 | 0.68605 | 1.6955 | 0.73052 | 1.2671 | 0.94868 | 0.30605 | 2.8062 | 0.58572 | 0.81296
| Rel.Std.Dev. |  | 2.853% | 2.007% | 3.398% | 19.73% | 9.219% | 5.005% | 3.222% | 1.991% | 3.533% | 1.801% | 4.152% | 2.361% | 3.923% | 3.108% | 1.909% | 4.844% | 2.368% | 5.428% | 2.154% | 4.064% | 1.972% | 0.8497% | 7.442% | 2.271% | 2.728%
| PDM3 |  | 2.962% | -2.141% | 2.241% | 125.3% | -29.16% | -16.23% | 13.17% | -2.696% | -35.65% | -4.582% | -22.41% | 22.92% | 9.398% | -10.24% | -4.693% | -8.355% | -3.584% | 3.961% | 12.89% | 3.795% | 60.1% | 19.88% | 25.5% | -14.14% | -0.8023%
| Z-Score (Absolute) |  | 0.25 | 0.09 | 0.2 | 8.34 | 1.87 | 1.02 | 0.92 | 0.13 | 2.3 | 0.25 | 1.43 | 1.57 | 0.67 | 0.62 | 0.26 | 0.5 | 0.18 | 0.31 | 0.91 | 0.3 | 4.03 | 1.37 | 1.74 | 0.88 | 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C74: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C92: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C129: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C149: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C548: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C786: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C987: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1151: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1187: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1205: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1241: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 45f (Certified Value 0.318 wt.%) | [28 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 37 | 39 | 40 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.3 | 0.33 | 0.31 | 0.32 | 0.35115 | 0.3 | 0.28 | 0.33 | 0.31 | 0.322 | 0.3 | 0.31 | 0.33 | 0.32 | 0.3435 | 0.334 | 0.34 | 0.29 | 0.33 | 0.34 | 0.32 | 0.308 | 0.32 | 0.34 | 0.28 | 0.312 | 0.31487 | 1
| 1 | 2 | 0.4 | 0.33 | 0.307 | 0.31 | 0.36 | 0.29 | 0.26 | 0.33 | 0.3 | 0.327 | 0.3 | 0.31 | 0.32 | 0.32 | 0.35033 | 0.308 | 0.34 | 0.28 | 0.32 | 0.34 | 0.33 | 0.299 | 0.32 | 0.35 | 0.27 | 0.315 | 0.30886 | #N/A
| 1 | 3 | 0.4 | 0.32 | 0.304 | 0.31 | 0.358 | 0.29 | 0.26 | 0.33 | 0.29 | 0.33 | 0.3 | 0.3 | 0.32 | 0.32 | 0.3325 | 0.303 | 0.35 | 0.29 | 0.34 | 0.36 | 0.33 | 0.302 | 0.33 | 0.35 | 0.27 | 0.307 | 0.31599 | 16
| 1 | 4 | 0.4 | 0.32 | 0.304 | 0.31 | 0.3615 | 0.29 | 0.26 | 0.33 | 0.33 | 0.333 | 0.3 | 0.31 | 0.32 | 0.32 | 0.341 | 0.303 | 0.35 | 0.3 | 0.34 | 0.32 | 0.32 | 0.306 | 0.32 | 0.35 | 0.28 | 0.323 | 0.31416 | 0.31845
| 1 | 5 | 0.3 | 0.33 | 0.306 | 0.31 | 0.365 | 0.29 | 0.24 | 0.33 | NA | 0.337 | 0.3 | 0.31 | 0.32 | 0.32 | 0.336 | 0.3 | 0.34 | 0.28 | 0.34 | 0.34 | 0.33 | 0.308 | 0.32 | 0.36 | 0.27 | 0.318 | 0.31438 | 192
| 1 | 6 | 0.3 | 0.33 | 0.303 | 0.32 | 0.3595 | 0.29 | 0.25 | 0.33 | 0.32 | 0.334 | 0.3 | 0.31 | 0.33 | 0.32 | 0.3475 | 0.32 | 0.35 | 0.28 | 0.34 | 0.34 | 0.33 | 0.308 | 0.32 | 0.35 | 0.28 | 0.313 | 0.30728
| Mean |  | 0.35 | 0.32667 | 0.30567 | 0.31333 | 0.35919 | 0.29167 | 0.25833 | 0.33 | 0.31 | 0.3305 | 0.3 | 0.30833 | 0.32333 | 0.32 | 0.34181 | 0.31133 | 0.345 | 0.28667 | 0.335 | 0.34 | 0.32667 | 0.30517 | 0.32167 | 0.35 | 0.275 | 0.31467 | 0.31259
| Median |  | 0.35 | 0.33 | 0.305 | 0.31 | 0.35975 | 0.29 | 0.26 | 0.33 | 0.31 | 0.3315 | 0.3 | 0.31 | 0.32 | 0.32 | 0.34225 | 0.3055 | 0.345 | 0.285 | 0.34 | 0.34 | 0.33 | 0.307 | 0.32 | 0.35 | 0.275 | 0.314 | 0.31427
| Std Dev. |  | 0.054772 | 0.005164 | 0.002582 | 0.005164 | 0.0046022 | 0.0040825 | 0.013292 | 0 | 0.015811 | 0.0053944 | 0 | 0.0040825 | 0.005164 | 0 | 0.0067671 | 0.013171 | 0.0054772 | 0.008165 | 0.0083666 | 0.012649 | 0.005164 | 0.0038166 | 0.0040825 | 0.0063246 | 0.0054772 | 0.005465 | 0.0035927
| Rel.Std.Dev. |  | 15.65% | 1.581% | 0.8447% | 1.648% | 1.281% | 1.4% | 5.145% | 0% | 5.1% | 1.632% | 0% | 1.324% | 1.597% | 0% | 1.98% | 4.23% | 1.588% | 2.848% | 2.497% | 3.72% | 1.581% | 1.251% | 1.269% | 1.807% | 1.992% | 1.737% | 1.149%
| PDM3 |  | 9.907% | 2.58% | -4.014% | -1.607% | 12.79% | -8.411% | -18.88% | 3.627% | -2.654% | 3.784% | -5.794% | -3.177% | 1.533% | 0.4865% | 7.334% | -2.235% | 8.337% | -9.981% | 5.197% | 6.767% | 2.58% | -4.171% | 1.01% | 9.907% | -13.64% | -1.188% | -1.84%
| Z-Score (Absolute) |  | 1.41 | 0.31 | 0.67 | 0.31 | 1.84 | 1.33 | 2.9 | 0.47 | 0.47 | 0.49 | 0.94 | 0.55 | 0.16 | 0 | 1.03 | 0.41 | 1.18 | 1.57 | 0.71 | 0.94 | 0.31 | 0.7 | 0.08 | 1.41 | 2.12 | 0.25 | 0.35
| Analytical results for S in OREAS 45f (Indicative Value 0.025 wt.%) | [28 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 04 | 05 | 07 | 10 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 25 | 27 | 28 | 29 | 30 | 31 | 33 | 34 | 35 | 37 | 39 | 40 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | < 0.1 | 0.03 | 0.018 | 0.03 | 0.019365 | 0.02 | 0.03 | 0.03 | < 0.01 | 0.014 | 0.03 | < 0.02 | 0.04 | 0.04 | 0.029 | 0.0281 | 0.02 | 0.03 | 0.02 | 0.02 | 0.05 | < 0.005 | < 0.01 | 0.02 | 0.06 | 0.018 | 0.0043419 | 1
| 1 | 2 | < 0.1 | 0.02 | 0.014 | 0.03 | 0.01895 | 0.03 | 0.03 | 0.03 | < 0.01 | 0.013 | 0.03 | 0.02 | 0.02 | 0.03 | 0.029 | 0.0291 | 0.02 | 0.03 | 0.01 | 0.02 | 0.05 | < 0.005 | < 0.01 | 0.03 | 0.07 | 0.02 | 0.0058539 | 28
| 1 | 3 | < 0.1 | 0.03 | 0.02 | 0.03 | 0.019 | 0.04 | 0.04 | 0.02 | < 0.01 | 0.009 | 0.06 | < 0.02 | 0.02 | 0.03 | 0.028 | 0.0322 | 0.01 | 0.03 | 0.02 | 0.04 | 0.05 | < 0.005 | 0.02 | 0.03 | 0.06 | 0.02 | < 0.003 | 16
| 1 | 4 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.016855 | 0.03 | 0.03 | 0.04 | < 0.01 | 0.014 | 0.06 | < 0.02 | < 0.01 | 0.03 | 0.027585 | 0.0275 | 0.02 | 0.04 | 0.02 | 0.05 | 0.05 | < 0.005 | < 0.01 | 0.02 | 0.06 | 0.019 | 0.0076343 | 0.024832
| 1 | 5 | < 0.1 | 0.03 | 0.018 | 0.02 | 0.0185 | 0.02 | 0.03 | 0.03 | NA | 0.023 | 0.02 | < 0.02 | 0.01 | 0.03 | 0.03 | 0.0255 | 0.02 | 0.03 | 0.02 | 0.03 | 0.05 | < 0.005 | < 0.01 | 0.02 | 0.05 | 0.019 | < 0.003 | 34
| 1 | 6 | < 0.1 | 0.03 | 0.018 | 0.03 | 0.0174 | 0.03 | 0.03 | 0.03 | < 0.01 | 0.015 | 0.03 | < 0.02 | 0.01 | 0.03 | 0.029 | 0.0295 | 0.02 | 0.03 | 0.02 | 0.03 | 0.05 | < 0.005 | < 0.01 | 0.03 | 0.07 | 0.019 | 0.0063034
| Mean |  |  | 0.026667 | 0.018 | 0.028333 | 0.018345 | 0.028333 | 0.031667 | 0.03 |  | 0.014667 | 0.038333 | 0.02 | 0.02 | 0.031667 | 0.028764 | 0.02865 | 0.018333 | 0.031667 | 0.018333 | 0.031667 | 0.05 |  | 0.02 | 0.025 | 0.061667 | 0.019167 | 0.0060334
| Median |  |  | 0.03 | 0.018 | 0.03 | 0.018725 | 0.03 | 0.03 | 0.03 |  | 0.014 | 0.03 | 0.02 | 0.02 | 0.03 | 0.029 | 0.0286 | 0.02 | 0.03 | 0.02 | 0.03 | 0.05 |  | 0.02 | 0.025 | 0.06 | 0.019 | 0.0060786
| Std Dev. |  |  | 0.005164 | 0.0021909 | 0.0040825 | 0.00099726 | 0.0075277 | 0.0040825 | 0.0063246 |  | 0.0045898 | 0.017224 |  | 0.012247 | 0.0040825 | 0.00085653 | 0.0022394 | 0.0040825 | 0.0040825 | 0.0040825 | 0.01169 | 7.6012e-18 |  |  | 0.0054772 | 0.0075277 | 0.00075277 | 0.0013576
| Rel.Std.Dev. |  |  | 19.36% | 12.17% | 14.41% | 5.436% | 26.57% | 12.89% | 21.08% |  | 31.29% | 44.93% |  | 61.24% | 12.89% | 2.978% | 7.816% | 22.27% | 12.89% | 22.27% | 36.92% | 1.52e-14% |  |  | 21.91% | 12.21% | 3.928% | 22.5%
| PDM3 |  |  | 7.39% | -27.51% | 14.1% | -26.12% | 14.1% | 27.53% | 20.81% |  | -40.94% | 54.37% | -19.46% | -19.46% | 27.53% | 15.84% | 15.38% | -26.17% | 27.53% | -26.17% | 27.53% | 101.4% |  | -19.46% | 0.6783% | 148.3% | -22.81% | -75.7%
| Z-Score (Absolute) |  | 1.89 | 0 | 0.7 | 0.13 | 0.67 | 0.13 | 0.4 | 0.27 | 1.75 | 0.97 | 0.94 | 1.21 | 0.74 | 0.4 | 0.17 | 0.16 | 0.67 | 0.4 | 0.67 | 0.4 | 1.89 | 1.96 | 1.55 | 0.13 | 2.83 | 0.61 | 1.79
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.