## Abbreviations
Table 1. Abbreviations used for OREAS 24b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
2A*OES/MS | 2:1 HNO3 and HCL (for organic or high sulphide content) digest with ICP-OES or ICP-MS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
AR*GFAAS | aqua regia digest with graphite furnace AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
FA*MS | fire assay with ICP-MS finish
FA*OES | fire assay with ICP-OES finish
IRC | infrared combustion furnace
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 24b
Lab Slot | Laboratory Name and Location
Acme (BV), Vancouver, BC, Canada
Acme-2 (BV), Vancouver, BC, Canada
Actlabs, Ancaster, Ontario, Canada
Actlabs-2, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Vancouver, BC, Canada
Amdel (BV), Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Genalysis-2, Perth, WA, Australia
SGS Mineral Services, Toronto, Ontario, Canada
SGS Mineral Services-2, Toronto, Ontario, Canada
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
Ultra Trace Pty Ltd (BV), Perth, WA, Australia

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 24b
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | < 3 | IND | IND | IND | IND | IND
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 15.1457 | 0.094145 | 15.0887 | 15.2027 | 15.0657 | 15.2258
BaO, Barium oxide (ppm) | 819.475 | 45.3988 | 780.2622 | 858.6878 | IND | IND
CaO, Calcium oxide (wt.%) | 1.4728 | 0.022522 | 1.457 | 1.4885 | 1.4557 | 1.4898
Cr2O3, Chromium(III) oxide (ppm) | 201.3156 | 7.261 | 199.2833 | 203.348 | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 6.3498 | 0.058074 | 6.313 | 6.3866 | 6.3045 | 6.3952
K2O, Potassium oxide (wt.%) | 3.3919 | 0.02581 | 3.373 | 3.4108 | 3.3727 | 3.4111
MgO, Magnesium oxide (wt.%) | 2.7548 | 0.029237 | 2.7307 | 2.7788 | 2.736 | 2.7735
MnO, Manganese oxide (wt.%) | 0.058905 | 0.0018439 | 0.05721 | 0.0606 | IND | IND
Na2O, Sodium oxide (wt.%) | 1.1494 | 0.023725 | 1.131 | 1.1678 | 1.1301 | 1.1686
P2O5, Phosphorus(V) oxide (wt.%) | 0.16077 | 0.0053731 | 0.15604 | 0.1655 | 0.15811 | 0.16343
SiO2, Silicon dioxide (wt.%) | 65.9996 | 0.26339 | 65.8491 | 66.1501 | 65.7523 | 66.247
TiO2, Titanium dioxide (wt.%) | 0.79831 | 0.018677 | 0.7844 | 0.81223 | 0.78736 | 0.80927
Thermogravimetry
LOI1000, Loss on ignition @1000°C (wt.%) | 2.4615 | 0.16897 | 2.3372 | 2.5859 | 2.4243 | 2.4988
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 7.8113 | 0.18046 | 7.7341 | 7.8884 | 7.6848 | 7.9378
Ba, Barium (ppm) | 738.9374 | 29.4858 | 718.7871 | 759.0878 | 722.5904 | 755.2844
Be, Beryllium (ppm) | 2.9461 | 0.37369 | 2.8223 | 3.07 | IND | IND
Ca, Calcium (wt.%) | 1.0562 | 0.041415 | 1.03 | 1.0823 | 1.0352 | 1.0771
Ce, Cerium (ppm) | 85.9039 | 2.4991 | 84.4798 | 87.3281 | 83.7978 | 88.0101
Co, Cobalt (ppm) | 16.9403 | 0.82695 | 16.3071 | 17.5735 | 16.5062 | 17.3743
Cr, Chromium (ppm) | 141.7297 | 12.8404 | 132.835 | 150.6245 | 130.4065 | 153.053
Cs, Cesium (ppm) | 10.4884 | 0.42743 | 10.2188 | 10.7579 | 10.173 | 10.8038
Dy, Dysprosium (ppm) | 5.8344 | 0.24988 | 5.7128 | 5.956 | 5.5372 | 6.1316
Er, Erbium (ppm) | 3.4128 | 0.12974 | 3.3505 | 3.4752 | 3.264 | 3.5617
Eu, Europium (ppm) | 1.3938 | 0.086111 | 1.346 | 1.4416 | 1.3458 | 1.4418
Fe, Iron (wt.%) | 4.4546 | 0.124 | 4.3851 | 4.5241 | 4.389 | 4.5202
Ga, Gallium (ppm) | 20.1584 | 0.87381 | 19.5634 | 20.7534 | 19.5474 | 20.7694
Gd, Gadolinium (ppm) | 6.2658 | 0.37409 | 5.9733 | 6.5584 | 6.0044 | 6.5272
Hf, Hafnium (ppm) | 6.1473 | 0.46134 | 5.8412 | 6.4535 | 5.8803 | 6.4144
Ho, Holmium (ppm) | 1.1715 | 0.063396 | 1.1355 | 1.2075 | 1.135 | 1.2079
K, Potassium (wt.%) | 2.7401 | 0.051259 | 2.716 | 2.7643 | 2.689 | 2.7912
La, Lanthanum (ppm) | 44.0168 | 1.538 | 43.0874 | 44.9462 | 42.6992 | 45.3344
Li, Lithium (ppm) | 51.8011 | 3.1595 | 47.9758 | 55.6264 | 49.513 | 54.0893
Lu, Lutetium (ppm) | 0.48881 | 0.044682 | 0.45363 | 0.52399 | 0.4551 | 0.52252
Mg, Magnesium (wt.%) | 1.6185 | 0.042782 | 1.5949 | 1.6422 | 1.5948 | 1.6423
Mn, Manganese (wt.%) | 0.046175 | 0.00047238 | 0.045798 | 0.046552 | 0.045061 | 0.047289
Na, Sodium (wt.%) | 0.82354 | 0.022353 | 0.79511 | 0.85198 | 0.80452 | 0.84256
Nb, Niobium (ppm) | 15.9957 | 0.64668 | 15.54 | 16.4515 | 15.3824 | 16.6091
Nd, Neodymium (ppm) | 38.6759 | 0.79247 | 38.1882 | 39.1637 | 37.4627 | 39.8892
Ni, Nickel (ppm) | 61.4693 | 7.163 | 56.8285 | 66.11 | 55.0065 | 67.9321
Pr, Praseodymium (ppm) | 10.1851 | 0.36935 | 9.9388 | 10.4314 | 9.9422 | 10.428
Rb, Rubidium (ppm) | 161.085 | 3.3786 | 158.4363 | 163.7337 | 157.6105 | 164.5595
S, Sulphur (wt.%) | 0.20333 | 0.034324 | 0.1671 | 0.23957 | IND | IND
Sc, Scandium (ppm) | 14.1229 | 0.71436 | 13.2457 | 15.0002 | IND | IND
Si, Silicon (wt.%) | 31.1209 | 1.0549 | 30.3523 | 31.8894 | 30.7177 | 31.524
Sm, Samarium (ppm) | 7.1738 | 0.28306 | 7.0253 | 7.3223 | 6.8367 | 7.5109
Sn, Tin (ppm) | 4.6535 | 0.67704 | 4.1829 | 5.1241 | IND | IND
Sr, Strontium (ppm) | 124.6251 | 4.0321 | 121.9264 | 127.3239 | 121.7759 | 127.4744
Ta, Tantalum (ppm) | 1.3197 | 0.083225 | 1.2675 | 1.3719 | IND | IND
Tb, Terbium (ppm) | 0.97969 | 0.045468 | 0.95151 | 1.0079 | 0.93792 | 1.0215
Th, Thorium (ppm) | 16.5355 | 0.64944 | 16.1401 | 16.9309 | 16.1651 | 16.9059
Ti, Titanium (wt.%) | 0.48057 | 0.017063 | 0.47061 | 0.49053 | 0.47065 | 0.49049
Tl, Thallium (ppm) | 0.90532 | 0.095803 | 0.806 | 1.0046 | IND | IND
Tm, Thulium (ppm) | 0.50377 | 0.029373 | 0.48373 | 0.52381 | 0.47899 | 0.52855
U, Uranium (ppm) | 3.3137 | 0.13555 | 3.231 | 3.3965 | 3.1841 | 3.4434
V, Vanadium (ppm) | 112.0198 | 10.4503 | 104.6085 | 119.431 | 108.6659 | 115.3736
W, Tungsten (ppm) | 4.1331 | 0.52113 | 3.8206 | 4.4456 | 1.7373 | 6.529
Y, Yttrium (ppm) | 32.5042 | 1.5269 | 31.4642 | 33.5442 | 31.6216 | 33.3868
Yb, Ytterbium (ppm) | 3.24 | 0.13477 | 3.1704 | 3.3096 | 3.0951 | 3.385
Zr, Zirconium (ppm) | 213.455 | 10.5281 | 206.4097 | 220.5003 | 206.6065 | 220.3034
4-Acid Digestion
Al, Aluminium (wt.%) | 8.0164 | 0.48406 | 7.5509 | 8.482 | 7.6962 | 8.3366
Ba, Barium (ppm) | 715.7515 | 32.5905 | 689.7815 | 741.7214 | 697.7427 | 733.7602
Be, Beryllium (ppm) | 2.9242 | 0.47351 | 2.5304 | 3.3179 | 2.7302 | 3.1181
Bi, Bismuth (ppm) | 0.6776 | 0.063237 | 0.63605 | 0.71914 | IND | IND
Ca, Calcium (wt.%) | 1.0838 | 0.040272 | 1.0428 | 1.1247 | 1.0534 | 1.1142
Ce, Cerium (ppm) | 83.6617 | 3.492 | 80.6136 | 86.7098 | 81.7509 | 85.5725
Co, Cobalt (ppm) | 16.876 | 1.1871 | 16.3105 | 17.4415 | 15.9718 | 17.7803
Cr, Chromium (ppm) | 118.2905 | 9.9115 | 109.4529 | 127.1281 | 111.5406 | 125.0403
Cs, Cesium (ppm) | 10.6533 | 0.40654 | 10.252 | 11.0546 | 10.2609 | 11.0457
Cu, Copper (ppm) | 37.9595 | 1.5875 | 36.8691 | 39.05 | 35.7516 | 40.1675
Fe, Iron (wt.%) | 4.3941 | 0.13121 | 4.2699 | 4.5182 | 4.2771 | 4.511
Ga, Gallium (ppm) | 20.1212 | 0.81632 | 19.5228 | 20.7197 | 19.2357 | 21.0068
Hf, Hafnium (ppm) | 3.9008 | 0.54021 | 3.4355 | 4.366 | 3.7303 | 4.0712
K, Potassium (wt.%) | 2.8062 | 0.10987 | 2.7214 | 2.8911 | 2.7171 | 2.8954
La, Lanthanum (ppm) | 42.3548 | 2.3356 | 40.4222 | 44.2874 | 40.7948 | 43.9147
Li, Lithium (ppm) | 51.7003 | 1.1674 | 50.968 | 52.4326 | 49.9233 | 53.4773
Mg, Magnesium (wt.%) | 1.6479 | 0.053594 | 1.594 | 1.7019 | 1.5934 | 1.7025
Mn, Manganese (wt.%) | 0.044226 | 0.0020997 | 0.042673 | 0.045779 | 0.042959 | 0.045493
Mo, Molybdenum (ppm) | 4.0324 | 0.35685 | 3.7296 | 4.3352 | 3.7103 | 4.3546
Na, Sodium (wt.%) | 0.84574 | 0.025233 | 0.82531 | 0.86617 | 0.81913 | 0.87234
Nb, Niobium (ppm) | 14.5979 | 1.2851 | 13.4311 | 15.7647 | 14.0544 | 15.1414
Ni, Nickel (ppm) | 60.0776 | 4.1364 | 56.6201 | 63.5351 | 57.4579 | 62.6973
P, Phosphorus (wt.%) | 0.069287 | 0.0049787 | 0.065053 | 0.073521 | 0.067539 | 0.071036
Pb, Lead (ppm) | 23.1472 | 1.384 | 22.0711 | 24.2234 | 21.9467 | 24.3477
Rb, Rubidium (ppm) | 163.6294 | 8.5351 | 156.4966 | 170.7623 | 156.5216 | 170.7373
S, Sulphur (wt.%) | 0.1983 | 0.0082417 | 0.19058 | 0.20603 | 0.18912 | 0.20749
Sb, Antimony (ppm) | 0.99725 | 0.045085 | 0.97067 | 1.0238 | IND | IND
Sc, Scandium (ppm) | 15.2548 | 1.4329 | 14.0782 | 16.4314 | IND | IND
Sn, Tin (ppm) | 4.2509 | 0.22062 | 4.0563 | 4.4454 | 4.0608 | 4.4409
Sr, Strontium (ppm) | 123.8498 | 11.3714 | 114.975 | 132.7246 | 120.2627 | 127.437
Ta, Tantalum (ppm) | 1.2273 | 0.21987 | 1.0219 | 1.4327 | 1.1597 | 1.2948
Th, Thorium (ppm) | 16.421 | 1.1058 | 15.5242 | 17.3179 | 15.772 | 17.0701
Ti, Titanium (wt.%) | 0.46836 | 0.017601 | 0.44655 | 0.49016 | 0.45775 | 0.47896
Tl, Thallium (ppm) | 0.86207 | 0.042546 | 0.81383 | 0.91032 | 0.80563 | 0.91852
U, Uranium (ppm) | 3.0593 | 0.093271 | 2.985 | 3.1336 | 2.9317 | 3.1869
V, Vanadium (ppm) | 108.0611 | 4.3719 | 104.138 | 111.9843 | 104.1791 | 111.9432
W, Tungsten (ppm) | 3.6391 | 0.47497 | 3.2703 | 4.0079 | 3.0671 | 4.2112
Y, Yttrium (ppm) | 19.8712 | 2.0413 | 17.9292 | 21.8132 | 18.8159 | 20.9264
Zn, Zinc (ppm) | 104.6524 | 5.6599 | 100.4944 | 108.8104 | 100.7934 | 108.5113
Zr, Zirconium (ppm) | 134.4032 | 24.0232 | 113.1209 | 155.6856 | 130.2477 | 138.5587
Aqua Regia Digestion (sample weights 0.15-50g)
Al, Aluminium (wt.%) | 3.152 | 0.15073 | 3.0168 | 3.2872 | 3.0528 | 3.2512
As, Arsenic (ppm) | 7.9567 | 0.74047 | 7.641 | 8.2724 | 7.4812 | 8.4323
Ba, Barium (ppm) | 145.7952 | 6.1422 | 139.5908 | 151.9997 | 139.8618 | 151.7287
Bi, Bismuth (ppm) | 0.73357 | 0.070646 | 0.67548 | 0.79167 | 0.67152 | 0.79562
Ca, Calcium (wt.%) | 0.46078 | 0.017547 | 0.44641 | 0.47516 | 0.44567 | 0.47589
Co, Cobalt (ppm) | 15.6533 | 1.3076 | 14.6589 | 16.6478 | 14.6577 | 16.6489
Cr, Chromium (ppm) | 106.4188 | 2.7591 | 104.8198 | 108.0177 | 102.2471 | 110.5904
Cs, Cesium (ppm) | 9.1493 | 0.60957 | 8.3791 | 9.9195 | 8.8666 | 9.4321
Cu, Copper (ppm) | 36.3669 | 1.3103 | 35.3989 | 37.335 | 34.6903 | 38.0436
Fe, Iron (wt.%) | 3.9257 | 0.25398 | 3.7118 | 4.1395 | 3.8249 | 4.0265
Ga, Gallium (ppm) | 10.765 | 1.108 | 9.8219 | 11.7081 | 10.2132 | 11.3168
K, Potassium (wt.%) | 1.1733 | 0.035002 | 1.1436 | 1.2029 | 1.1361 | 1.2104
Li, Lithium (ppm) | 45.6248 | 6.252 | 37.4926 | 53.757 | 43.8936 | 47.356
Mg, Magnesium (wt.%) | 1.3646 | 0.071378 | 1.3059 | 1.4233 | 1.3271 | 1.4021
Mn, Manganese (wt.%) | 0.034992 | 0.0015474 | 0.033604 | 0.03638 | 0.033846 | 0.036138
Mo, Molybdenum (ppm) | 3.8606 | 0.22023 | 3.6769 | 4.0443 | 3.675 | 4.0462
Na, Sodium (wt.%) | 0.10783 | 0.0094564 | 0.099359 | 0.11631 | IND | IND
Ni, Nickel (ppm) | 56.5306 | 4.1271 | 53.3815 | 59.6797 | 53.6864 | 59.3748
Pb, Lead (ppm) | 9.2339 | 1.0066 | 8.5627 | 9.9051 | 8.5926 | 9.8751
Rb, Rubidium (ppm) | 114.3427 | 10.3873 | 101.1092 | 127.5762 | 106.737 | 121.9484
S, Sulphur (wt.%) | 0.19992 | 0.0098239 | 0.18972 | 0.21012 | 0.18958 | 0.21026
Sc, Scandium (ppm) | 9.5128 | 0.71796 | 8.9451 | 10.0804 | 8.8165 | 10.209
Sn, Tin (ppm) | 2.2602 | 0.2078 | 2.0369 | 2.4835 | 2.1067 | 2.4137
Th, Thorium (ppm) | 14.3325 | 2.2931 | 11.78 | 16.8849 | 13.795 | 14.8699
Ti, Titanium (wt.%) | 0.19849 | 0.010569 | 0.18558 | 0.21141 | 0.19013 | 0.20685
Tl, Thallium (ppm) | 0.65777 | 0.081368 | 0.56305 | 0.75249 | 0.61966 | 0.69588
U, Uranium (ppm) | 1.7417 | 0.32016 | 1.3837 | 2.0996 | 1.6109 | 1.8725
V, Vanadium (ppm) | 78.9213 | 1.7051 | 77.5935 | 80.2491 | 76.4285 | 81.414
Y, Yttrium (ppm) | 12.3449 | 1.8221 | 10.3074 | 14.3823 | 11.5305 | 13.1592
Zn, Zinc (ppm) | 92.7511 | 4.0536 | 89.4688 | 96.0333 | 89.8645 | 95.6376
Zr, Zirconium (ppm) | 24.467 | 3.6142 | 19.7485 | 29.1855 | 22.8682 | 26.0658
Infrared Combustion
C, Carbon (wt.%) | 0.18876 | 0.011214 | 0.17902 | 0.19849 | IND | IND
S, Sulphur (wt.%) | 0.19006 | 0.013408 | 0.17929 | 0.20084 | IND | IND

## Indicative Values
Table 3. Indicative Values for OREAS 24b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 1 | Pt | ppb | < 1
Borate Fusion XRF
As | ppm | 10 | Ni | ppm | 58.8425 | V2O5 | ppm | 184.1667
Cl | ppm | < 10 | S | wt.% | 0.19544 | Zn | ppm | 113.3333
Co | ppm | 28.3333 | Sr | ppm | 133.7563
Borate / Peroxide Fusion ICP
Ag | ppm | 2.173 | Ge | ppm | 1.6399 | Sb | ppm | 1.3261
As | ppm | 9.7352 | In | ppm | < 0.2 | Se | ppm | < 10
B | ppm | 69.1667 | Mo | ppm | 4.9107 | Te | ppm | < 6
Bi | ppm | 1.0258 | P | wt.% | 0.072933 | Zn | ppm | 102.69
Cd | ppm | < 2 | Pb | ppm | 22.894
Cu | ppm | 35.1403 | Re | ppm | < 0.1
4-Acid Digestion
Ag | ppm | 0.12684 | Ge | ppm | 0.83444 | Se | ppm | 0.66185
As | ppm | 8.35 | Ho | ppm | 0.80111 | Sm | ppm | 7.0583
Cd | ppm | 0.049312 | In | ppm | 0.076718 | Tb | ppm | 0.8655
Dy | ppm | 4.4656 | Lu | ppm | 0.32258 | Te | ppm | < 0.1
Er | ppm | 2.5383 | Nd | ppm | 36.2389 | Tm | ppm | 0.30667
Eu | ppm | 1.3612 | Pr | ppm | 9.86 | Yb | ppm | 2.1659
Gd | ppm | 6.0244 | Re | ppm | 0.0015833
Aqua Regia Digestion (sample weights 0.15-50g)
Ag | ppm | 0.057508 | Hg | ppm | < 0.01 | Sb | ppm | 0.47841
Au | ppm | 0.0023407 | Ho | ppm | 0.45833 | Se | ppm | 0.4192
B | ppm | 6.2269 | In | ppm | 0.047627 | Sm | ppm | 4.6808
Be | ppm | 1.6547 | La | ppm | 29.1621 | Sr | ppm | 29.0079
Cd | ppm | 0.045875 | Lu | ppm | 0.19566 | Ta | ppm | < 0.05
Ce | ppm | 61.2619 | Nb | ppm | 0.30833 | Tb | ppm | 0.53775
Dy | ppm | 2.6475 | Nd | ppm | 24.6333 | Te | ppm | < 0.02
Er | ppm | 1.2083 | P | wt.% | 0.062229 | Tm | ppm | 0.17292
Eu | ppm | 0.65833 | Pd | ppm | < 0.01 | W | ppm | 1.1886
Gd | ppm | 3.9633 | Pr | ppm | 6.8658 | Yb | ppm | 1.1505
Ge | ppm | 0.25665 | Pt | ppm | 0.0013333
Hf | ppm | 0.52292 | Re | ppm | < 0.001

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 24b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | < 3 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Borate Fusion XRF
Al2O3, wt.% | 15.1457 | 0.094145 | 14.9575 | 15.334 | 14.8633 | 15.4282 | 0.6216% | 1.243% | 1.865% | 14.3885 | 15.903
BaO, ppm | 819.475 | 45.3988 | 728.6775 | 910.2725 | 683.2787 | 955.6713 | 5.54% | 11.08% | 16.62% | 778.5013 | 860.4488
CaO, wt.% | 1.4728 | 0.022522 | 1.4277 | 1.5178 | 1.4052 | 1.5403 | 1.529% | 3.058% | 4.588% | 1.3991 | 1.5464
Cr2O3, ppm | 201.3156 | 7.261 | 186.7936 | 215.8377 | 179.5326 | 223.0987 | 3.607% | 7.214% | 10.82% | 191.2499 | 211.3814
Fe2O3, wt.% | 6.3498 | 0.058074 | 6.2337 | 6.466 | 6.1756 | 6.5241 | 0.9146% | 1.829% | 2.744% | 6.0323 | 6.6673
K2O, wt.% | 3.3919 | 0.02581 | 3.3403 | 3.4435 | 3.3145 | 3.4693 | 0.7609% | 1.522% | 2.283% | 3.2223 | 3.5615
MgO, wt.% | 2.7548 | 0.029237 | 2.6963 | 2.8132 | 2.667 | 2.8425 | 1.061% | 2.123% | 3.184% | 2.617 | 2.8925
MnO, wt.% | 0.058905 | 0.0018439 | 0.055217 | 0.062593 | 0.053373 | 0.064436 | 3.13% | 6.261% | 9.391% | 0.05596 | 0.06185
Na2O, wt.% | 1.1494 | 0.023725 | 1.1019 | 1.1968 | 1.0782 | 1.2206 | 2.064% | 4.128% | 6.192% | 1.0919 | 1.2069
P2O5, wt.% | 0.16077 | 0.0053731 | 0.15002 | 0.17152 | 0.14465 | 0.17689 | 3.342% | 6.684% | 10.03% | 0.15273 | 0.16881
SiO2, wt.% | 65.9996 | 0.26339 | 65.4728 | 66.5264 | 65.2095 | 66.7898 | 0.3991% | 0.7982% | 1.197% | 62.6996 | 69.2996
TiO2, wt.% | 0.79831 | 0.018677 | 0.76096 | 0.83567 | 0.74228 | 0.85434 | 2.34% | 4.679% | 7.019% | 0.7584 | 0.83823
Thermogravimetry
LOI1000, wt.% | 2.4615 | 0.16897 | 2.1236 | 2.7995 | 1.9546 | 2.9684 | 6.864% | 13.73% | 20.59% | 2.3385 | 2.5846
Borate / Peroxide Fusion ICP
Al, wt.% | 7.8113 | 0.18046 | 7.4503 | 8.1722 | 7.2699 | 8.3526 | 2.31% | 4.62% | 6.931% | 7.4207 | 8.2018
Ba, ppm | 738.9374 | 29.4858 | 679.9659 | 797.9089 | 650.4802 | 827.3947 | 3.99% | 7.981% | 11.97% | 701.9906 | 775.8843
Be, ppm | 2.9461 | 0.37369 | 2.1987 | 3.6935 | 1.8251 | 4.0672 | 12.68% | 25.37% | 38.05% | 2.7988 | 3.0934
Ca, wt.% | 1.0562 | 0.041415 | 0.97333 | 1.139 | 0.93192 | 1.1804 | 3.921% | 7.843% | 11.76% | 1.0034 | 1.109
Ce, ppm | 85.9039 | 2.4991 | 80.9057 | 90.9022 | 78.4066 | 93.4013 | 2.909% | 5.818% | 8.728% | 81.6087 | 90.1991
Co, ppm | 16.9403 | 0.82695 | 15.2863 | 18.5942 | 14.4594 | 19.4211 | 4.882% | 9.763% | 14.64% | 16.0932 | 17.7873
Cr, ppm | 141.7297 | 12.8404 | 116.0489 | 167.4106 | 103.2085 | 180.251 | 9.06% | 18.12% | 27.18% | 134.6432 | 148.8162
Cs, ppm | 10.4884 | 0.42743 | 9.6335 | 11.3432 | 9.2061 | 11.7706 | 4.075% | 8.15% | 12.23% | 9.9639 | 11.0128
Dy, ppm | 5.8344 | 0.24988 | 5.3346 | 6.3342 | 5.0847 | 6.584 | 4.283% | 8.566% | 12.85% | 5.5427 | 6.1261
Er, ppm | 3.4128 | 0.12974 | 3.1533 | 3.6723 | 3.0236 | 3.802 | 3.802% | 7.603% | 11.4% | 3.2422 | 3.5835
Eu, ppm | 1.3938 | 0.086111 | 1.2216 | 1.566 | 1.1354 | 1.6521 | 6.178% | 12.36% | 18.53% | 1.3241 | 1.4635
Fe, wt.% | 4.4546 | 0.124 | 4.2066 | 4.7026 | 4.0826 | 4.8266 | 2.784% | 5.567% | 8.351% | 4.2319 | 4.6774
Ga, ppm | 20.1584 | 0.87381 | 18.4108 | 21.906 | 17.537 | 22.7798 | 4.335% | 8.669% | 13% | 19.1505 | 21.1663
Gd, ppm | 6.2658 | 0.37409 | 5.5176 | 7.014 | 5.1435 | 7.3881 | 5.97% | 11.94% | 17.91% | 5.9525 | 6.5791
Hf, ppm | 6.1473 | 0.46134 | 5.2247 | 7.07 | 4.7633 | 7.5313 | 7.505% | 15.01% | 22.51% | 5.84 | 6.4547
Ho, ppm | 1.1715 | 0.063396 | 1.0447 | 1.2983 | 0.98127 | 1.3616 | 5.412% | 10.82% | 16.24% | 1.1129 | 1.23
K, wt.% | 2.7401 | 0.051259 | 2.6376 | 2.8426 | 2.5863 | 2.8939 | 1.871% | 3.741% | 5.612% | 2.6031 | 2.8771
La, ppm | 44.0168 | 1.538 | 40.9407 | 47.0929 | 39.4027 | 48.631 | 3.494% | 6.988% | 10.48% | 41.816 | 46.2177
Li, ppm | 51.8011 | 3.1595 | 45.4821 | 58.1202 | 42.3225 | 61.2798 | 6.099% | 12.2% | 18.3% | 49.2111 | 54.3912
Lu, ppm | 0.48881 | 0.044682 | 0.39944 | 0.57817 | 0.35476 | 0.62286 | 9.141% | 18.28% | 27.42% | 0.46437 | 0.51325
Mg, wt.% | 1.6185 | 0.042782 | 1.533 | 1.7041 | 1.4902 | 1.7469 | 2.643% | 5.286% | 7.93% | 1.5376 | 1.6995
Mn, wt.% | 0.046175 | 0.00047238 | 0.04523 | 0.04712 | 0.044758 | 0.047592 | 1.023% | 2.046% | 3.069% | 0.043866 | 0.048484
Na, wt.% | 0.82354 | 0.022353 | 0.77884 | 0.86825 | 0.75649 | 0.8906 | 2.714% | 5.428% | 8.143% | 0.78237 | 0.86472
Nb, ppm | 15.9957 | 0.64668 | 14.7024 | 17.2891 | 14.0557 | 17.9358 | 4.043% | 8.086% | 12.13% | 15.196 | 16.7955
Nd, ppm | 38.6759 | 0.79247 | 37.091 | 40.2609 | 36.2985 | 41.0534 | 2.049% | 4.098% | 6.147% | 36.7422 | 40.6097
Ni, ppm | 61.4693 | 7.163 | 47.1434 | 75.7952 | 39.9804 | 82.9581 | 11.65% | 23.31% | 34.96% | 58.3958 | 64.5427
Pr, ppm | 10.1851 | 0.36935 | 9.4464 | 10.9238 | 9.077 | 11.2932 | 3.626% | 7.253% | 10.88% | 9.6758 | 10.6943
Rb, ppm | 161.085 | 3.3786 | 154.3278 | 167.8422 | 150.9493 | 171.2207 | 2.097% | 4.195% | 6.292% | 153.0308 | 169.1393
S, wt.% | 0.20333 | 0.034324 | 0.13468 | 0.27198 | 0.10036 | 0.30631 | 16.88% | 33.76% | 50.64% | 0.19317 | 0.2135
Sc, ppm | 14.1229 | 0.71436 | 12.6942 | 15.5516 | 11.9798 | 16.266 | 5.058% | 10.12% | 15.17% | 13.4168 | 14.8291
Si, wt.% | 31.1209 | 1.0549 | 29.0111 | 33.2306 | 27.9563 | 34.2854 | 3.39% | 6.779% | 10.17% | 29.5648 | 32.6769
Sm, ppm | 7.1738 | 0.28306 | 6.6077 | 7.7399 | 6.3246 | 8.023 | 3.946% | 7.892% | 11.84% | 6.8151 | 7.5325
Sn, ppm | 4.6535 | 0.67704 | 3.2994 | 6.0076 | 2.6224 | 6.6846 | 14.55% | 29.1% | 43.65% | 4.4208 | 4.8862
Sr, ppm | 124.6251 | 4.0321 | 116.561 | 132.6893 | 112.5289 | 136.7213 | 3.235% | 6.471% | 9.706% | 118.3939 | 130.8564
Ta, ppm | 1.3197 | 0.083225 | 1.1533 | 1.4862 | 1.07 | 1.5694 | 6.306% | 12.61% | 18.92% | 1.2537 | 1.3857
Tb, ppm | 0.97969 | 0.045468 | 0.88875 | 1.0706 | 0.84328 | 1.1161 | 4.641% | 9.282% | 13.92% | 0.9307 | 1.0287
Th, ppm | 16.5355 | 0.64944 | 15.2366 | 17.8344 | 14.5872 | 18.4838 | 3.928% | 7.855% | 11.78% | 15.7087 | 17.3623
Ti, wt.% | 0.48057 | 0.017063 | 0.44644 | 0.51469 | 0.42938 | 0.53176 | 3.551% | 7.101% | 10.65% | 0.45654 | 0.5046
Tl, ppm | 0.90532 | 0.095803 | 0.71371 | 1.0969 | 0.61791 | 1.1927 | 10.58% | 21.16% | 31.75% | 0.86005 | 0.95058
Tm, ppm | 0.50377 | 0.029373 | 0.44502 | 0.56251 | 0.41565 | 0.59189 | 5.831% | 11.66% | 17.49% | 0.47858 | 0.52896
U, ppm | 3.3137 | 0.13555 | 3.0426 | 3.5849 | 2.9071 | 3.7204 | 4.091% | 8.181% | 12.27% | 3.1481 | 3.4794
V, ppm | 112.0198 | 10.4503 | 91.1192 | 132.9203 | 80.669 | 143.3706 | 9.329% | 18.66% | 27.99% | 106.4188 | 117.6208
W, ppm | 4.1331 | 0.52113 | 3.0908 | 5.1754 | 2.5697 | 5.6965 | 12.61% | 25.22% | 37.83% | 3.9265 | 4.3398
Y, ppm | 32.5042 | 1.5269 | 29.4504 | 35.558 | 27.9235 | 37.0849 | 4.698% | 9.395% | 14.09% | 30.879 | 34.1294
Yb, ppm | 3.24 | 0.13477 | 2.9705 | 3.5096 | 2.8357 | 3.6443 | 4.16% | 8.319% | 12.48% | 3.078 | 3.402
Zr, ppm | 213.455 | 10.5281 | 192.3987 | 234.5113 | 181.8705 | 245.0394 | 4.932% | 9.865% | 14.8% | 202.7822 | 224.1277
4-Acid Digestion
Al, wt.% | 8.0164 | 0.48406 | 7.0483 | 8.9845 | 6.5642 | 9.4686 | 6.038% | 12.08% | 18.11% | 7.6156 | 8.4172
Ba, ppm | 715.7515 | 32.5905 | 650.5704 | 780.9325 | 617.9799 | 813.5231 | 4.553% | 9.107% | 13.66% | 679.9639 | 751.539
Be, ppm | 2.9242 | 0.47351 | 1.9772 | 3.8712 | 1.5036 | 4.3447 | 16.19% | 32.39% | 48.58% | 2.778 | 3.0704
Bi, ppm | 0.6776 | 0.063237 | 0.55112 | 0.80407 | 0.48789 | 0.86731 | 9.333% | 18.67% | 28% | 0.64372 | 0.71147
Ca, wt.% | 1.0838 | 0.040272 | 1.0032 | 1.1643 | 0.96294 | 1.2046 | 3.716% | 7.432% | 11.15% | 1.0296 | 1.1379
Ce, ppm | 83.6617 | 3.492 | 76.6777 | 90.6456 | 73.1857 | 94.1376 | 4.174% | 8.348% | 12.52% | 79.4786 | 87.8448
Co, ppm | 16.876 | 1.1871 | 14.5019 | 19.2501 | 13.3148 | 20.4372 | 7.034% | 14.07% | 21.1% | 16.0322 | 17.7198
Cr, ppm | 118.2905 | 9.9115 | 98.4675 | 138.1135 | 88.556 | 148.025 | 8.379% | 16.76% | 25.14% | 112.376 | 124.205
Cs, ppm | 10.6533 | 0.40654 | 9.8402 | 11.4664 | 9.4336 | 11.8729 | 3.816% | 7.632% | 11.45% | 10.1206 | 11.1859
Cu, ppm | 37.9595 | 1.5875 | 34.7846 | 41.1344 | 33.1972 | 42.7219 | 4.182% | 8.364% | 12.55% | 36.0616 | 39.8575
Fe, wt.% | 4.3941 | 0.13121 | 4.1316 | 4.6565 | 4.0004 | 4.7877 | 2.986% | 5.972% | 8.959% | 4.1743 | 4.6138
Ga, ppm | 20.1212 | 0.81632 | 18.4886 | 21.7539 | 17.6723 | 22.5702 | 4.057% | 8.114% | 12.17% | 19.1152 | 21.1273
Hf, ppm | 3.9008 | 0.54021 | 2.8204 | 4.9812 | 2.2801 | 5.5214 | 13.85% | 27.7% | 41.55% | 3.7057 | 4.0958
K, wt.% | 2.8062 | 0.10987 | 2.5865 | 3.026 | 2.4766 | 3.1359 | 3.915% | 7.831% | 11.75% | 2.6659 | 2.9466
La, ppm | 42.3548 | 2.3356 | 37.6836 | 47.026 | 35.3479 | 49.3616 | 5.514% | 11.03% | 16.54% | 40.2371 | 44.4725
Li, ppm | 51.7003 | 1.1674 | 49.3656 | 54.035 | 48.1982 | 55.2023 | 2.258% | 4.516% | 6.774% | 49.1153 | 54.2853
Mg, wt.% | 1.6479 | 0.053594 | 1.5407 | 1.7551 | 1.4872 | 1.8087 | 3.252% | 6.504% | 9.756% | 1.5655 | 1.7303
Mn, wt.% | 0.044226 | 0.0020997 | 0.040027 | 0.048426 | 0.037927 | 0.050525 | 4.748% | 9.495% | 14.24% | 0.042015 | 0.046437
Mo, ppm | 4.0324 | 0.35685 | 3.3187 | 4.7461 | 2.9619 | 5.103 | 8.85% | 17.7% | 26.55% | 3.8308 | 4.2341
Na, wt.% | 0.84574 | 0.025233 | 0.79527 | 0.8962 | 0.77004 | 0.92144 | 2.984% | 5.967% | 8.951% | 0.80345 | 0.88802
Nb, ppm | 14.5979 | 1.2851 | 12.0276 | 17.1682 | 10.7425 | 18.4533 | 8.804% | 17.61% | 26.41% | 13.868 | 15.3278
Ni, ppm | 60.0776 | 4.1364 | 51.8048 | 68.3504 | 47.6683 | 72.4868 | 6.885% | 13.77% | 20.66% | 57.0737 | 63.0815
P, wt.% | 0.069287 | 0.0049787 | 0.05933 | 0.079245 | 0.054351 | 0.084223 | 7.186% | 14.37% | 21.56% | 0.065823 | 0.072752
Pb, ppm | 23.1472 | 1.384 | 20.3793 | 25.9152 | 18.9953 | 27.2992 | 5.979% | 11.96% | 17.94% | 21.9899 | 24.3046
Rb, ppm | 163.6294 | 8.5351 | 146.5593 | 180.6995 | 138.0242 | 189.2346 | 5.216% | 10.43% | 15.65% | 155.4479 | 171.8109
S, wt.% | 0.1983 | 0.0082417 | 0.18182 | 0.21479 | 0.17358 | 0.22303 | 4.156% | 8.312% | 12.47% | 0.18839 | 0.20822
Sb, ppm | 0.99725 | 0.045085 | 0.90708 | 1.0874 | 0.862 | 1.1325 | 4.521% | 9.042% | 13.56% | 0.94739 | 1.0471
Sc, ppm | 15.2548 | 1.4329 | 12.3891 | 18.1205 | 10.9562 | 19.5534 | 9.393% | 18.79% | 28.18% | 14.492 | 16.0175
Sn, ppm | 4.2509 | 0.22062 | 3.8096 | 4.6921 | 3.589 | 4.9127 | 5.19% | 10.38% | 15.57% | 4.0383 | 4.4634
Sr, ppm | 123.8498 | 11.3714 | 101.1069 | 146.5927 | 89.7355 | 157.9641 | 9.182% | 18.36% | 27.54% | 117.6573 | 130.0423
Ta, ppm | 1.2273 | 0.21987 | 0.78754 | 1.667 | 0.56767 | 1.8869 | 17.92% | 35.83% | 53.75% | 1.1659 | 1.2887
Th, ppm | 16.421 | 1.1058 | 14.2095 | 18.6326 | 13.1037 | 19.7384 | 6.734% | 13.47% | 20.2% | 15.6 | 17.2421
Ti, wt.% | 0.46836 | 0.017601 | 0.43315 | 0.50356 | 0.41555 | 0.52116 | 3.758% | 7.516% | 11.27% | 0.44494 | 0.49178
Tl, ppm | 0.86207 | 0.042546 | 0.77698 | 0.94716 | 0.73443 | 0.98971 | 4.935% | 9.871% | 14.81% | 0.81897 | 0.90517
U, ppm | 3.0593 | 0.093271 | 2.8727 | 3.2458 | 2.7795 | 3.3391 | 3.049% | 6.098% | 9.146% | 2.9063 | 3.2123
V, ppm | 108.0611 | 4.3719 | 99.3173 | 116.805 | 94.9454 | 121.1769 | 4.046% | 8.092% | 12.14% | 102.6581 | 113.4642
W, ppm | 3.6391 | 0.47497 | 2.6892 | 4.5891 | 2.2142 | 5.0641 | 13.05% | 26.1% | 39.16% | 3.4572 | 3.8211
Y, ppm | 19.8712 | 2.0413 | 15.7886 | 23.9538 | 13.7473 | 25.9951 | 10.27% | 20.55% | 30.82% | 18.8776 | 20.8648
Zn, ppm | 104.6524 | 5.6599 | 93.3325 | 115.9722 | 87.6726 | 121.6322 | 5.408% | 10.82% | 16.22% | 99.4198 | 109.885
Zr, ppm | 134.4032 | 24.0232 | 86.3569 | 182.4496 | 62.3337 | 206.4728 | 17.87% | 35.75% | 53.62% | 127.6831 | 141.1234
Aqua Regia Digestion (sample weights 0.15-50g)
Al, wt.% | 3.152 | 0.15073 | 2.8506 | 3.4535 | 2.6998 | 3.6042 | 4.782% | 9.564% | 14.35% | 2.9944 | 3.3096
As, ppm | 7.9567 | 0.74047 | 6.4758 | 9.4377 | 5.7353 | 10.1781 | 9.306% | 18.61% | 27.92% | 7.5589 | 8.3546
Ba, ppm | 145.7952 | 6.1422 | 133.5108 | 158.0797 | 127.3686 | 164.2219 | 4.213% | 8.426% | 12.64% | 138.5055 | 153.085
Bi, ppm | 0.73357 | 0.070646 | 0.59228 | 0.87487 | 0.52164 | 0.94551 | 9.63% | 19.26% | 28.89% | 0.69689 | 0.77025
Ca, wt.% | 0.46078 | 0.017547 | 0.42569 | 0.49588 | 0.40814 | 0.51342 | 3.808% | 7.616% | 11.42% | 0.43774 | 0.48382
Co, ppm | 15.6533 | 1.3076 | 13.0382 | 18.2685 | 11.7306 | 19.5761 | 8.353% | 16.71% | 25.06% | 14.8707 | 16.436
Cr, ppm | 106.4188 | 2.7591 | 100.9006 | 111.9369 | 98.1415 | 114.696 | 2.593% | 5.185% | 7.778% | 101.0978 | 111.7397
Cs, ppm | 9.1493 | 0.60957 | 7.9302 | 10.3685 | 7.3206 | 10.978 | 6.662% | 13.32% | 19.99% | 8.6919 | 9.6068
Cu, ppm | 36.3669 | 1.3103 | 33.7463 | 38.9876 | 32.436 | 40.2979 | 3.603% | 7.206% | 10.81% | 34.5486 | 38.1853
Fe, wt.% | 3.9257 | 0.25398 | 3.4177 | 4.4336 | 3.1637 | 4.6876 | 6.47% | 12.94% | 19.41% | 3.7294 | 4.122
Ga, ppm | 10.765 | 1.108 | 8.549 | 12.981 | 7.441 | 14.089 | 10.29% | 20.59% | 30.88% | 10.2267 | 11.3033
K, wt.% | 1.1733 | 0.035002 | 1.1033 | 1.2433 | 1.0682 | 1.2783 | 2.983% | 5.967% | 8.95% | 1.1146 | 1.2319
Li, ppm | 45.6248 | 6.252 | 33.1208 | 58.1289 | 26.8687 | 64.3809 | 13.7% | 27.41% | 41.11% | 43.3436 | 47.9061
Mg, wt.% | 1.3646 | 0.071378 | 1.2218 | 1.5073 | 1.1504 | 1.5787 | 5.231% | 10.46% | 15.69% | 1.2963 | 1.4328
Mn, wt.% | 0.034992 | 0.0015474 | 0.031897 | 0.038087 | 0.03035 | 0.039634 | 4.422% | 8.844% | 13.27% | 0.033242 | 0.036742
Mo, ppm | 3.8606 | 0.22023 | 3.4201 | 4.3011 | 3.1999 | 4.5213 | 5.705% | 11.41% | 17.11% | 3.6676 | 4.0536
Na, wt.% | 0.10783 | 0.0094564 | 0.088921 | 0.12675 | 0.079464 | 0.1362 | 8.769% | 17.54% | 26.31% | 0.10244 | 0.11323
Ni, ppm | 56.5306 | 4.1271 | 48.2764 | 64.7847 | 44.1494 | 68.9118 | 7.301% | 14.6% | 21.9% | 53.704 | 59.3571
Pb, ppm | 9.2339 | 1.0066 | 7.2207 | 11.2471 | 6.2141 | 12.2536 | 10.9% | 21.8% | 32.7% | 8.7722 | 9.6956
Rb, ppm | 114.3427 | 10.3873 | 93.5681 | 135.1173 | 83.1808 | 145.5046 | 9.084% | 18.17% | 27.25% | 108.6256 | 120.0598
S, wt.% | 0.19992 | 0.0098239 | 0.18027 | 0.21957 | 0.17045 | 0.22939 | 4.914% | 9.828% | 14.74% | 0.18992 | 0.20992
Sc, ppm | 9.5128 | 0.71796 | 8.0768 | 10.9487 | 7.3589 | 11.6666 | 7.547% | 15.09% | 22.64% | 9.0371 | 9.9884
Sn, ppm | 2.2602 | 0.2078 | 1.8446 | 2.6758 | 1.6368 | 2.8836 | 9.194% | 18.39% | 27.58% | 2.1472 | 2.3732
Th, ppm | 14.3325 | 2.2931 | 9.7462 | 18.9187 | 7.4531 | 21.2118 | 16% | 32% | 48% | 13.6158 | 15.0491
Ti, wt.% | 0.19849 | 0.010569 | 0.17735 | 0.21963 | 0.16678 | 0.2302 | 5.325% | 10.65% | 15.97% | 0.18857 | 0.20842
Tl, ppm | 0.65777 | 0.081368 | 0.49503 | 0.82051 | 0.41367 | 0.90188 | 12.37% | 24.74% | 37.11% | 0.62488 | 0.69066
U, ppm | 1.7417 | 0.32016 | 1.1014 | 2.382 | 0.78119 | 2.7021 | 18.38% | 36.76% | 55.15% | 1.6546 | 1.8288
V, ppm | 78.9213 | 1.7051 | 75.5111 | 82.3314 | 73.8061 | 84.0365 | 2.16% | 4.321% | 6.481% | 74.9752 | 82.8673
Y, ppm | 12.3449 | 1.8221 | 8.7006 | 15.9891 | 6.8785 | 17.8112 | 14.76% | 29.52% | 44.28% | 11.7276 | 12.9621
Zn, ppm | 92.7511 | 4.0536 | 84.6439 | 100.8582 | 80.5903 | 104.9118 | 4.37% | 8.741% | 13.11% | 88.1135 | 97.3886
Zr, ppm | 24.467 | 3.6142 | 17.2386 | 31.6954 | 13.6244 | 35.3096 | 14.77% | 29.54% | 44.32% | 23.2436 | 25.6903
Infrared Combustion
C, wt.% | 0.18876 | 0.011214 | 0.16633 | 0.21118 | 0.15512 | 0.2224 | 5.941% | 11.88% | 17.82% | 0.17932 | 0.19819
S, wt.% | 0.19006 | 0.013408 | 0.16325 | 0.21688 | 0.14984 | 0.23029 | 7.055% | 14.11% | 21.16% | 0.18056 | 0.19957

## Fire Assay
| Analytical results for Au in OREAS 24b (Certified Value < 3 ppb) | [10 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | ppb
|  |  | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*MS | FA*MS | FA*MS | FA*OES | 2
|  |  | 30g | 40g | 40g | 25g | 30g | 30g | 30g | 30g | 30g | 2
| 1 | 1 | 2 | 2 | 1.8995 | < 1 | 1.69 | < 2 | 4 | 2 | 1 | 1
| 1 | 2 | 2 | 1 | 2.3102 | 3 | 1.56 | < 2 | 2 | < 1 | 5 | 23
| 1 | 3 | 3 | 2 | 1.8103 | 2 | 1.76 | < 2 | 2 | < 1 | < 1 | 16
| 1 | 4 | 3 | 4 | 2.0652 | 2 | 1.42 | < 2 | 2 | < 1 | 2 | < 3
| 1 | 5 | 3 | 2 | 1.9452 | 3 | 1.34 | < 2 | 2 | < 1 | 1 | 7
| 1 | 6 | 2 | 2 | 1.7905 | 2 | 1.45 | < 2 | 3 | < 1 | 6
| Mean |  | 2.5 | 2.1667 | 1.9702 | 2.4 | 1.5367 |  | 2.5 | 2 | 3
| Median |  | 2.5 | 2 | 1.9224 | 2 | 1.505 |  | 2 | 2 | 2
| Std Dev. |  | 0.54772 | 0.98319 | 0.19396 | 0.54772 | 0.16354 |  | 0.83666 |  | 2.3452
| Rel.Std.Dev. |  | 21.91% | 45.38% | 9.845% | 22.82% | 10.64% |  | 33.47% |  | 78.17%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.13 | 0.18 | 0 | 0.88 | 1.75 | 0.67 | 2.16 | 0.81
| Analytical results for Pd in OREAS 24b (Indicative Value < 1 ppb) | [10 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 |  | ppb
|  |  | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*MS | FA*MS | FA*OES |  | 2
|  |  | 30g | 40g | 40g | 25g | 30g | 30g | 30g | 30g |  | 2
| 1 | 1 | 1 | < 1 | < 1 | < 1 | 0.78 | < 1 | < 0.5 | < 1 |  | 1
| 1 | 2 | 1 | < 1 | < 1 | < 1 | 0.64 | < 1 | < 0.5 | < 1 |  | 1
| 1 | 3 | 1 | < 1 | < 1 | < 1 | 0.64 | < 1 | < 0.5 | < 1 |  | 16
| 1 | 4 | 1 | < 1 | 1.0994 | < 1 | 0.94 | < 1 | < 0.5 | < 1 |  | < 1
| 1 | 5 | 2 | 1 | < 1 | < 1 | 0.58 | < 1 | < 0.5 | < 1 |  | 7
| 1 | 6 | 1 | < 1 | < 1 | < 1 | 0.69 | < 1 | < 0.5 | < 1
| Mean |  | 01: 1.1667; 02: 1; 03: 1.0994; 06: 0.71167
| Median |  | 01: 1; 02: 1; 03: 1.0994; 06: 0.665
| Std Dev. |  | 01: 0.40825; 06: 0.13029
| Rel.Std.Dev. |  | 01: 34.99%; 06: 18.31%
PDM3
| Z-Score (Absolute) |  | 8.44 | 0.56 | 0.79 | 0.56 | 2.29 | 0.56 | 3.94 | 0.56
| Analytical results for Pt in OREAS 24b (Indicative Value < 1 ppb) | [10 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 |  | ppb
|  |  | FA*OES | FA*MS | FA*OES | FA*MS | FA*OES | FA*MS | FA*MS | FA*OES |  | 2
|  |  | 30g | 40g | 40g | 25g | 30g | 30g | 30g | 30g |  | 2
| 1 | 1 | < 5 | < 1 | < 5 | < 1 | 0.96 | < 1 | 0.6 | < 5 |  | 1
| 1 | 2 | < 5 | < 1 | < 5 | < 1 | 1.68 | < 1 | 0.5 | < 5 |  | 1
| 1 | 3 | < 5 | < 1 | < 5 | < 1 | 1.77 | < 1 | 0.4 | < 5 |  | 16
| 1 | 4 | 5 | < 1 | < 5 | < 1 | 0.69 | < 1 | 0.5 | < 5 |  | < 1
| 1 | 5 | < 5 | < 1 | < 5 | < 1 | 2.14 | < 1 | 0.4 | < 5 |  | 7
| 1 | 6 | < 5 | < 1 | < 5 | < 1 | 1.93 | < 1 | 0.5 | < 5
| Mean |  | 01: 5; 06: 1.5283; 09: 0.48333
| Median |  | 01: 5; 06: 1.725; 09: 0.5
| Std Dev. |  | 06: 0.57318; 09: 0.075277
| Rel.Std.Dev. |  | 06: 37.5%; 09: 15.57%
PDM3
| Z-Score (Absolute) |  | 2.46 | 0.66 | 1.92 | 0.66 | 0.66 | 0.66 | 0.69 | 1.92
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 24b (Certified Value 15.15 wt.%) | [10 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 15.17 | 15.07 | 15.27 | 15.2 | 15.15 | 15.15 | 15.07 | 15.24 | 15.01 | 1
| 1 | 2 | 15.15 | 15.09 | 15.21 | 14.8 | 15.2 | 15.15 | 15.11 | 15.25 | 15.06 | 0
| 1 | 3 | 15.11 | 15.1 | 15.27 | 15.2 | 15.15 | 15.25 | 15.1 | 15.18 | 15.07 | 16
| 1 | 4 | 15.12 | 15.08 | 15.26 | 15.2 | 15.15 | 15.1 | 15.02 | 15.23 | 14.98 | 15.1457
| 1 | 5 | 15.18 | 15.08 | 15.32 | 14.9 | 15.2 | 15.3 | 15.17 | 15.03 | 14.94 | 9
| 1 | 6 | 15.19 | 15.06 | 15.22 | 15.2 | 15.2 | 15.35 | 15.09 | 15.15 | 15.03
| Mean |  | 15.1533 | 15.08 | 15.2583 | 15.0833 | 15.175 | 15.2167 | 15.0933 | 15.18 | 15.015
| Median |  | 15.16 | 15.08 | 15.265 | 15.2 | 15.175 | 15.2 | 15.095 | 15.205 | 15.02
| Std Dev. |  | 0.03266 | 0.014142 | 0.039707 | 0.18348 | 0.027386 | 0.098319 | 0.049261 | 0.082946 | 0.049295
| Rel.Std.Dev. |  | 0.2155% | 0.09378% | 0.2602% | 1.216% | 0.1805% | 0.6461% | 0.3264% | 0.5464% | 0.3283%
| PDM3 |  | 0.05013% | -0.4341% | 0.7434% | -0.412% | 0.1932% | 0.4683% | -0.346% | 0.2262% | -0.8632%
| Z-Score (Absolute) |  | 0 | 0.78 | 1.12 | 0.75 | 0.23 | 0.67 | 0.64 | 0.28 | 1.47
| Analytical results for As in OREAS 24b (Indicative Value 10 ppm) | [10 empty] | Indicative
As | Round | Replicate | Lab | [8 empty] | 1
| No | No | 03 | [8 empty] | ppm
|  |  | BF*XRF | [8 empty] | 1
1
| 1 | 1 | 10 | [8 empty] | 1
| 1 | 2 | 10 | [8 empty] | 3
| 1 | 3 | 10 | [8 empty] | 16
| 1 | 4 | 20 | [8 empty] | 10
| 1 | 5 | 10 | [8 empty] | 9
| 1 | 6 | 10
| Mean |  | 11.6667
| Median |  | 10
| Std Dev. |  | 4.0825
| Rel.Std.Dev. |  | 34.99%
| PDM3 |  | 16.67%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 24b (Certified Value 819 ppm) | [10 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 08 | 10 | 11 | 14 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 900 | 692 | 800 | 770 | 800 | 890 |  |  |  | 1
| 1 | 2 | 800 | 681 | 800 | 790 | 700 | 890 |  |  |  | 0
| 1 | 3 | 900 | 715 | 800 | 820 | 800 | 780 |  |  |  | 16
| 1 | 4 | 900 | 703 | 800 | 770 | 800 | 890 |  |  |  | 819.475
| 1 | 5 | 800 | 670 | 800 | 810 | 800 | 780 |  |  |  | 10
| 1 | 6 | 900 | 648 | 800 | 800 | 800 | 780
| Mean |  | 866.6667 | 684.8333 | 800 | 793.3333 | 783.3333 | 835
| Median |  | 900 | 686.5 | 800 | 795 | 800 | 835
| Std Dev. |  | 51.6398 | 24.0118 | 0 | 20.6559 | 40.8248 | 60.2495
| Rel.Std.Dev. |  | 5.958% | 3.506% | 0% | 2.604% | 5.212% | 7.216%
| PDM3 |  | 5.759% | -16.43% | -2.377% | -3.19% | -4.41% | 1.895%
| Z-Score (Absolute) |  | 1.75 | 2.79 | 0.08 | 0.08 | 0.33 | 1.02
| Analytical results for CaO in OREAS 24b (Certified Value 1.47 wt.%) | [10 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.43 | 1.47 | 1.49 | 1.51 | 1.46 | 1.46 | 1.49 | 1.43 | 1.47 | 1
| 1 | 2 | 1.44 | 1.47 | 1.49 | 1.47 | 1.48 | 1.47 | 1.5 | 1.44 | 1.48 | 0
| 1 | 3 | 1.44 | 1.47 | 1.5 | 1.51 | 1.47 | 1.46 | 1.49 | 1.45 | 1.49 | 16
| 1 | 4 | 1.44 | 1.48 | 1.49 | 1.53 | 1.47 | 1.46 | 1.48 | 1.43 | 1.49 | 1.4728
| 1 | 5 | 1.44 | 1.48 | 1.5 | 1.46 | 1.47 | 1.48 | 1.5 | 1.45 | 1.47 | 11
| 1 | 6 | 1.43 | 1.48 | 1.49 | 1.49 | 1.48 | 1.48 | 1.48 | 1.47 | 1.48
| Mean |  | 1.4367 | 1.475 | 1.4933 | 1.495 | 1.4717 | 1.4683 | 1.49 | 1.445 | 1.48
| Median |  | 1.44 | 1.475 | 1.49 | 1.5 | 1.47 | 1.465 | 1.49 | 1.445 | 1.48
| Std Dev. |  | 0.005164 | 0.0054772 | 0.005164 | 0.026646 | 0.0075277 | 0.0098319 | 0.0089443 | 0.015166 | 0.0089443
| Rel.Std.Dev. |  | 0.3594% | 0.3713% | 0.3458% | 1.782% | 0.5115% | 0.6696% | 0.6003% | 1.05% | 0.6043%
| PDM3 |  | -2.452% | 0.1509% | 1.396% | 1.509% | -0.07544% | -0.3018% | 1.169% | -1.886% | 0.4904%
| Z-Score (Absolute) |  | 1.72 | 0 | 0.82 | 0.9 | 0.15 | 0.3 | 0.67 | 1.35 | 0.22
| Analytical results for Cl in OREAS 24b (Indicative Value < 10 ppm) | [10 empty] | Indicative
Cl | Round | Replicate | Lab | [8 empty] | 1
| No | No | 03 | [8 empty] | ppm
|  |  | BF*XRF | [8 empty] | 1
1
| 1 | 1 | < 10 | [8 empty] | 1
| 1 | 2 | < 10 | [8 empty] | 4
| 1 | 3 | < 10 | [8 empty] | 16
| 1 | 4 | < 10 | [8 empty] | < 10
| 1 | 5 | < 10 | [8 empty] | 10
| 1 | 6 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 24b (Indicative Value 28.3 ppm) | [10 empty] | Indicative
Co | Round | Replicate | Lab | Lab | [7 empty] | 1
| No | No | 03 | 13 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | [7 empty] | 1
1
| 1 | 1 | 20 | < 68 | [7 empty] | 1
| 1 | 2 | 30 | < 68 | [7 empty] | 5
| 1 | 3 | 30 | < 68 | [7 empty] | 16
| 1 | 4 | 30 | < 68 | [7 empty] | 28.3333
| 1 | 5 | 30 | < 68 | [7 empty] | 11
| 1 | 6 | 30 | < 68
| Mean |  | 28.3333
| Median |  | 30
| Std Dev. |  | 4.0825
| Rel.Std.Dev. |  | 14.41%
| PDM3 |  | 1.11e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cr2O3 in OREAS 24b (Certified Value 201 ppm) | [10 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 200 | 200 | 219 | < 100 | 200 | 190 | 200 | 200 | 220 | 1
| 1 | 2 | 200 | 190 | 205 | 200 | 100 | 210 | 200 | 200 | 200 | 0
| 1 | 3 | 200 | 200 | 190 | 200 | 200 | 210 | 200 | 300 | 200 | 16
| 1 | 4 | 100 | 200 | 190 | 200 | 100 | 200 | 200 | 200 | 210 | 201.3156
| 1 | 5 | 200 | 200 | 190 | 200 | 200 | 210 | 200 | 200 | 180 | 12
| 1 | 6 | 200 | 210 | 219 | 200 | 200 | 210 | 200 | 200 | 200
| Mean |  | 183.3333 | 200 | 202.1667 | 200 | 166.6667 | 205 | 200 | 216.6667 | 201.6667
| Median |  | 200 | 200 | 197.5 | 200 | 200 | 210 | 200 | 200 | 200
| Std Dev. |  | 40.8248 | 6.3246 | 14.2747 | 0 | 51.6398 | 8.3666 | 0 | 40.8248 | 13.2916
| Rel.Std.Dev. |  | 22.27% | 3.162% | 7.061% | 0% | 30.98% | 4.081% | 0% | 18.84% | 6.591%
| PDM3 |  | -8.932% | -0.6535% | 0.4227% | -0.6535% | -17.21% | 1.83% | -0.6535% | 7.625% | 0.1744%
| Z-Score (Absolute) |  | 2.25 | 0 | 0.29 | 3.37 | 4.5 | 0.67 | 0 | 2.25 | 0.22
| Analytical results for Fe2O3 in OREAS 24b (Certified Value 6.35 wt.%) | [10 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 6.31 | 6.34 | 6.291 | 6.4 | 6.33 | 6.35 | 6.43 | 6.36 | 6.3 | 1
| 1 | 2 | 6.32 | 6.36 | 6.262 | 6.2 | 6.33 | 6.39 | 6.44 | 6.38 | 6.36 | 0
| 1 | 3 | 6.31 | 6.36 | 6.276 | 6.35 | 6.3 | 6.39 | 6.43 | 6.41 | 6.41 | 16
| 1 | 4 | 6.3 | 6.38 | 6.262 | 6.37 | 6.33 | 6.33 | 6.4 | 6.36 | 6.38 | 6.3498
| 1 | 5 | 6.29 | 6.37 | 6.305 | 6.23 | 6.33 | 6.45 | 6.44 | 6.4 | 6.29 | 13
| 1 | 6 | 6.25 | 6.38 | 6.305 | 6.36 | 6.33 | 6.42 | 6.42 | 6.45 | 6.37
| Mean |  | 6.2967 | 6.365 | 6.2835 | 6.3183 | 6.325 | 6.3883 | 6.4267 | 6.3933 | 6.3517
| Median |  | 6.305 | 6.365 | 6.2835 | 6.355 | 6.33 | 6.39 | 6.43 | 6.39 | 6.365
| Std Dev. |  | 0.025033 | 0.015166 | 0.019807 | 0.082321 | 0.012247 | 0.044008 | 0.015055 | 0.034448 | 0.047081
| Rel.Std.Dev. |  | 0.3976% | 0.2383% | 0.3152% | 1.303% | 0.1936% | 0.6889% | 0.2343% | 0.5388% | 0.7412%
| PDM3 |  | -0.8373% | 0.2388% | -1.045% | -0.4961% | -0.3911% | 0.6063% | 1.21% | 0.685% | 0.02882%
| Z-Score (Absolute) |  | 1.01 | 0.25 | 1.25 | 0.61 | 0.49 | 0.67 | 1.38 | 0.77 | 0
| Analytical results for K2O in OREAS 24b (Certified Value 3.39 wt.%) | [10 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 3.3 | 3.38 | 3.41 | 3.4 | 3.38 | 3.34 | 3.4 | 3.38 | 3.45 | 1
| 1 | 2 | 3.26 | 3.39 | 3.42 | 3.3 | 3.38 | 3.37 | 3.43 | 3.38 | 3.43 | 0
| 1 | 3 | 3.28 | 3.39 | 3.41 | 3.38 | 3.37 | 3.38 | 3.42 | 3.36 | 3.41 | 16
| 1 | 4 | 3.28 | 3.38 | 3.41 | 3.38 | 3.38 | 3.35 | 3.41 | 3.36 | 3.43 | 3.3919
| 1 | 5 | 3.26 | 3.39 | 3.42 | 3.32 | 3.38 | 3.39 | 3.43 | 3.36 | 3.41 | 14
| 1 | 6 | 3.26 | 3.39 | 3.42 | 3.38 | 3.39 | 3.4 | 3.41 | 3.38 | 3.41
| Mean |  | 3.2733 | 3.3867 | 3.415 | 3.36 | 3.38 | 3.3717 | 3.4167 | 3.37 | 3.4233
| Median |  | 3.27 | 3.39 | 3.415 | 3.38 | 3.38 | 3.375 | 3.415 | 3.37 | 3.42
| Std Dev. |  | 0.01633 | 0.005164 | 0.0054772 | 0.04 | 0.0063246 | 0.023166 | 0.012111 | 0.010954 | 0.01633
| Rel.Std.Dev. |  | 0.4989% | 0.1525% | 0.1604% | 1.19% | 0.1871% | 0.6871% | 0.3545% | 0.3251% | 0.477%
| PDM3 |  | -3.496% | -0.1548% | 0.6805% | -0.941% | -0.3513% | -0.597% | 0.7297% | -0.6461% | 0.9262%
| Z-Score (Absolute) |  | 3.6 | 0.22 | 1.18 | 0.67 | 0 | 0.28 | 1.24 | 0.34 | 1.46
| Analytical results for MgO in OREAS 24b (Certified Value 2.75 wt.%) | [10 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.64 | 2.75 | 2.8 | 2.78 | 2.75 | 2.71 | 2.76 | 2.72 | 2.85 | 1
| 1 | 2 | 2.66 | 2.75 | 2.79 | 2.71 | 2.76 | 2.72 | 2.77 | 2.74 | 2.85 | 0
| 1 | 3 | 2.64 | 2.75 | 2.8 | 2.77 | 2.75 | 2.73 | 2.75 | 2.73 | 2.86 | 16
| 1 | 4 | 2.66 | 2.75 | 2.8 | 2.78 | 2.76 | 2.7 | 2.76 | 2.76 | 2.86 | 2.7548
| 1 | 5 | 2.63 | 2.75 | 2.81 | 2.74 | 2.76 | 2.74 | 2.78 | 2.7 | 2.85 | 15
| 1 | 6 | 2.67 | 2.75 | 2.81 | 2.81 | 2.75 | 2.73 | 2.74 | 2.73 | 2.85
| Mean |  | 2.65 | 2.75 | 2.8017 | 2.765 | 2.755 | 2.7217 | 2.76 | 2.73 | 2.8533
| Median |  | 2.65 | 2.75 | 2.8 | 2.775 | 2.755 | 2.725 | 2.76 | 2.73 | 2.85
| Std Dev. |  | 0.015492 | 0 | 0.0075277 | 0.035071 | 0.0054772 | 0.01472 | 0.014142 | 0.02 | 0.005164
| Rel.Std.Dev. |  | 0.5846% | 0% | 0.2687% | 1.268% | 0.1988% | 0.5408% | 0.5124% | 0.7326% | 0.181%
| PDM3 |  | -3.803% | -0.1729% | 1.703% | 0.3717% | 0.008643% | -1.201% | 0.1901% | -0.8989% | 3.578%
| Z-Score (Absolute) |  | 2.83 | 0.13 | 1.26 | 0.27 | 0 | 0.9 | 0.13 | 0.67 | 2.65
| Analytical results for MnO in OREAS 24b (Certified Value 0.059 wt.%) | [10 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.05 | 0.058 | 0.07 | 0.06 | 0.06 | 0.058 | 0.06 | 0.053 | 0.06 | 1
| 1 | 2 | 0.05 | 0.059 | 0.07 | 0.06 | 0.06 | 0.06 | 0.06 | 0.052 | 0.06 | 0
| 1 | 3 | 0.05 | 0.059 | 0.07 | 0.06 | 0.06 | 0.059 | 0.06 | 0.055 | 0.06 | 16
| 1 | 4 | 0.05 | 0.058 | 0.07 | 0.06 | 0.06 | 0.058 | 0.06 | 0.053 | 0.06 | 0.058905
| 1 | 5 | 0.05 | 0.058 | 0.07 | 0.06 | 0.06 | 0.059 | 0.06 | 0.055 | 0.06 | 16
| 1 | 6 | 0.05 | 0.06 | 0.07 | 0.06 | 0.06 | 0.058 | 0.06 | 0.059 | 0.06
| Mean |  | 0.05 | 0.058667 | 0.07 | 0.06 | 0.06 | 0.058667 | 0.06 | 0.0545 | 0.06
| Median |  | 0.05 | 0.0585 | 0.07 | 0.06 | 0.06 | 0.0585 | 0.06 | 0.054 | 0.06
| Std Dev. |  | 7.6012e-18 | 0.0008165 | 0 | 0 | 0 | 0.0008165 | 0 | 0.00251 | 0
| Rel.Std.Dev. |  | 1.52e-14% | 1.392% | 0% | 0% | 0% | 1.392% | 0% | 4.605% | 0%
| PDM3 |  | -15.12% | -0.4042% | 18.84% | 1.859% | 1.859% | -0.4042% | 1.859% | -7.478% | 1.859%
| Z-Score (Absolute) |  | 5.06 | 0.67 | 5.06 | 0 | 0 | 0.67 | 0 | 2.78 | 0
| Analytical results for Na2O in OREAS 24b (Certified Value 1.15 wt.%) | [10 empty] | Certified
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.14 | 1.13 | 1.14 | 1.15 | 1.12 | 1.17 | 1.14 | 1.21 | 1.16 | 1
| 1 | 2 | 1.14 | 1.13 | 1.15 | 1.16 | 1.12 | 1.185 | 1.14 | 1.18 | 1.17 | 0
| 1 | 3 | 1.15 | 1.12 | 1.14 | 1.14 | 1.12 | 1.17 | 1.13 | 1.19 | 1.18 | 16
| 1 | 4 | 1.13 | 1.13 | 1.15 | 1.14 | 1.12 | 1.17 | 1.13 | 1.17 | 1.16 | 1.1494
| 1 | 5 | 1.12 | 1.13 | 1.16 | 1.14 | 1.12 | 1.185 | 1.12 | 1.27 | 1.17 | 17
| 1 | 6 | 1.13 | 1.13 | 1.14 | 1.15 | 1.12 | 1.175 | 1.14 | 1.21 | 1.16
| Mean |  | 1.135 | 1.1283 | 1.1467 | 1.1467 | 1.12 | 1.1758 | 1.1333 | 1.205 | 1.1667
| Median |  | 1.135 | 1.13 | 1.145 | 1.145 | 1.12 | 1.1725 | 1.135 | 1.2 | 1.165
| Std Dev. |  | 0.010488 | 0.0040825 | 0.008165 | 0.008165 | 0 | 0.0073598 | 0.008165 | 0.035637 | 0.008165
| Rel.Std.Dev. |  | 0.9241% | 0.3618% | 0.7121% | 0.7121% | 0% | 0.6259% | 0.7204% | 2.957% | 0.6999%
| PDM3 |  | -1.252% | -1.832% | -0.2368% | -0.2368% | -2.557% | 2.301% | -1.397% | 4.838% | 1.503%
| Z-Score (Absolute) |  | 0.43 | 0.67 | 0 | 0 | 0.98 | 1.07 | 0.49 | 2.15 | 0.74
| Analytical results for Ni in OREAS 24b (Indicative Value 59 ppm) | [10 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | [7 empty] | 1
| No | No | 03 | 13 | [7 empty] | ppm
|  |  | BF*XRF | BF*XRF | [7 empty] | 0
0
| 1 | 1 | 90 | 31 | [7 empty] | 1
| 1 | 2 | 60 | 31 | [7 empty] | 3
| 1 | 3 | 70 | 63 | [7 empty] | 16
| 1 | 4 | 80 | 39 | [7 empty] | 58.8425
| 1 | 5 | 50 | 39 | [7 empty] | 9
| 1 | 6 | 140 | < 24
| Mean |  | 81.6667 | 40.6
| Median |  | 75 | 39
| Std Dev. |  | 31.8852 | 13.1453
| Rel.Std.Dev. |  | 39.04% | 32.38%
| PDM3 |  | 38.79% | -31%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for P2O5 in OREAS 24b (Certified Value 0.161 wt.%) | [10 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.157 | 0.164 | 0.1421 | 0.16 | 0.15 | 0.16 | 0.162 | 0.17 | 0.16 | 1
| 1 | 2 | 0.158 | 0.162 | 0.1421 | 0.16 | 0.15 | 0.162 | 0.164 | 0.16 | 0.16 | 0
| 1 | 3 | 0.158 | 0.164 | 0.1421 | 0.16 | 0.15 | 0.162 | 0.163 | 0.17 | 0.16 | 16
| 1 | 4 | 0.158 | 0.164 | 0.1421 | 0.16 | 0.15 | 0.16 | 0.162 | 0.17 | 0.16 | 0.16077
| 1 | 5 | 0.159 | 0.165 | 0.1421 | 0.16 | 0.15 | 0.162 | 0.165 | 0.17 | 0.15 | 18
| 1 | 6 | 0.158 | 0.165 | 0.1398 | 0.16 | 0.15 | 0.164 | 0.159 | 0.17 | 0.16
| Mean |  | 0.158 | 0.164 | 0.14172 | 0.16 | 0.15 | 0.16167 | 0.1625 | 0.16833 | 0.15833
| Median |  | 0.158 | 0.164 | 0.1421 | 0.16 | 0.15 | 0.162 | 0.1625 | 0.17 | 0.16
| Std Dev. |  | 0.00063246 | 0.0010954 | 0.00093897 | 0 | 0 | 0.0015055 | 0.0020736 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 0.4003% | 0.668% | 0.6626% | 0% | 0% | 0.9313% | 1.276% | 2.425% | 2.578%
| PDM3 |  | -1.723% | 2.009% | -11.85% | -0.4795% | -6.699% | 0.5572% | 1.076% | 4.704% | -1.516%
| Z-Score (Absolute) |  | 0.54 | 1.08 | 4.94 | 0 | 2.7 | 0.45 | 0.67 | 2.25 | 0.45
| Analytical results for S in OREAS 24b (Indicative Value 0.195 wt.%) | [10 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 11 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.1922 | 0.195 | 0.2 | [6 empty] | 1
| 1 | 2 | 0.1922 | 0.195 | 0.2 | [6 empty] | 4
| 1 | 3 | 0.1922 | 0.194 | 0.2 | [6 empty] | 16
| 1 | 4 | 0.1918 | 0.195 | 0.2 | [6 empty] | 0.19544
| 1 | 5 | 0.1918 | 0.193 | 0.2 | [6 empty] | 10
| 1 | 6 | 0.1922 | 0.192 | 0.2
| Mean |  | 0.19207 | 0.194 | 0.2
| Median |  | 0.1922 | 0.1945 | 0.2
| Std Dev. |  | 0.00020656 | 0.0012649 | 3.0405e-17
| Rel.Std.Dev. |  | 0.1075% | 0.652% | 1.52e-14%
| PDM3 |  | -1.724% | -0.7344% | 2.336%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.19
| Analytical results for SiO2 in OREAS 24b (Certified Value 66 wt.%) | [10 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 65.96 | 66.16 | 65.74 | 66.5 | 65.7 | 66 | 66.06 | 65.81 | 66 | 1
| 1 | 2 | 65.99 | 66.12 | 65.52 | 66.5 | 65.9 | 66.3 | 66.18 | 66.1 | 66.1 | 0
| 1 | 3 | 65.67 | 66.12 | 65.72 | 66.1 | 65.7 | 66.4 | 65.99 | 65.74 | 66.2 | 16
| 1 | 4 | 65.6 | 66.14 | 65.74 | 66.4 | 65.8 | 65.9 | 65.84 | 65.6 | 66.1 | 65.9996
| 1 | 5 | 65.85 | 66.05 | 65.96 | 66.1 | 65.8 | 66.8 | 66.08 | 65.79 | 65.6 | 19
| 1 | 6 | 66.1 | 66.1 | 65.69 | 66 | 65.9 | 67 | 66.16 | 66.42 | 65.9
| Mean |  | 65.8617 | 66.115 | 65.7283 | 66.2667 | 65.8 | 66.4 | 66.0517 | 65.91 | 65.9833
| Median |  | 65.905 | 66.12 | 65.73 | 66.25 | 65.8 | 66.35 | 66.07 | 65.8 | 66.05
| Std Dev. |  | 0.19405 | 0.037815 | 0.14063 | 0.22509 | 0.089443 | 0.43359 | 0.12465 | 0.29853 | 0.2137
| Rel.Std.Dev. |  | 0.2946% | 0.0572% | 0.214% | 0.3397% | 0.1359% | 0.653% | 0.1887% | 0.4529% | 0.3239%
| PDM3 |  | -0.209% | 0.1748% | -0.4111% | 0.4046% | -0.3025% | 0.6066% | 0.07884% | -0.1358% | -0.02469%
| Z-Score (Absolute) |  | 0.62 | 0.67 | 1.31 | 1.45 | 0.94 | 2.13 | 0.35 | 0.38 | 0
| Analytical results for Sr in OREAS 24b (Indicative Value 134 ppm) | [10 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 08 | 10 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 170 | 110 | 170 | 110 | [5 empty] | 1
| 1 | 2 | 170 | 110 | 170 | 127 | [5 empty] | 5
| 1 | 3 | 170 | 110 | 170 | 118 | [5 empty] | 16
| 1 | 4 | 80 | 100 | 170 | 110 | [5 empty] | 133.7563
| 1 | 5 | 80 | 120 | 170 | 118 | [5 empty] | 11
| 1 | 6 | 170 | 110 | 170 | 118
| Mean |  | 140 | 110 | 170 | 116.8333
| Median |  | 170 | 110 | 170 | 118
| Std Dev. |  | 46.4758 | 6.3246 | 0 | 6.3377
| Rel.Std.Dev. |  | 33.2% | 5.75% | 0% | 5.425%
| PDM3 |  | 4.668% | -17.76% | 27.1% | -12.65%
| Z-Score (Absolute) |  | 0.52 | 0.83 | 1.75 | 0.52
| Analytical results for TiO2 in OREAS 24b (Certified Value 0.798 wt.%) | [10 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 08 | 10 | 11 | 13 | 14 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.79 | 0.797 | 0.8 | 0.82 | 0.79 | 0.79 | 0.81 | 0.76 | 0.85 | 1
| 1 | 2 | 0.78 | 0.796 | 0.8 | 0.78 | 0.79 | 0.81 | 0.81 | 0.76 | 0.83 | 0
| 1 | 3 | 0.78 | 0.799 | 0.8 | 0.81 | 0.79 | 0.81 | 0.81 | 0.76 | 0.8 | 16
| 1 | 4 | 0.78 | 0.798 | 0.8 | 0.81 | 0.79 | 0.79 | 0.8 | 0.76 | 0.82 | 0.79831
| 1 | 5 | 0.78 | 0.799 | 0.8 | 0.8 | 0.79 | 0.81 | 0.82 | 0.77 | 0.83 | 20
| 1 | 6 | 0.79 | 0.8 | 0.8 | 0.81 | 0.79 | 0.81 | 0.81 | 0.77 | 0.83
| Mean |  | 0.78333 | 0.79817 | 0.8 | 0.805 | 0.79 | 0.80333 | 0.81 | 0.76333 | 0.82667
| Median |  | 0.78 | 0.7985 | 0.8 | 0.81 | 0.79 | 0.81 | 0.81 | 0.76 | 0.83
| Std Dev. |  | 0.005164 | 0.001472 | 1.2162e-16 | 0.013784 | 0 | 0.010328 | 0.0063246 | 0.005164 | 0.01633
| Rel.Std.Dev. |  | 0.6592% | 0.1844% | 1.52e-14% | 1.712% | 0% | 1.286% | 0.7808% | 0.6765% | 1.975%
| PDM3 |  | -1.877% | -0.01856% | 0.2111% | 0.8374% | -1.042% | 0.6286% | 1.464% | -4.382% | 3.551%
| Z-Score (Absolute) |  | 1.12 | 0.12 | 0 | 0.34 | 0.67 | 0.22 | 0.67 | 2.47 | 1.8
| Analytical results for V2O5 in OREAS 24b (Indicative Value 184 ppm) | [10 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 13 | [6 empty] | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 214 | 300 | 190 | [6 empty] | 1
| 1 | 2 | 214 | < 100 | 170 | [6 empty] | 3
| 1 | 3 | 214 | 300 | 170 | [6 empty] | 16
| 1 | 4 | 196 | 200 | 200 | [6 empty] | 184.1667
| 1 | 5 | 214 | < 100 | 180 | [6 empty] | 9
| 1 | 6 | 214 | 200 | 200
| Mean |  | 211 | 250 | 185
| Median |  | 214 | 250 | 185
| Std Dev. |  | 7.3485 | 57.735 | 13.784
| Rel.Std.Dev. |  | 3.483% | 23.09% | 7.451%
| PDM3 |  | 14.57% | 35.75% | 0.4525%
| Z-Score (Absolute) |  | 10.62 | 0.67 | 0
| Analytical results for Zn in OREAS 24b (Indicative Value 113 ppm) | [10 empty] | Indicative
Zn | Round | Replicate | Lab | [8 empty] | 1
| No | No | 03 | [8 empty] | ppm
|  |  | BF*XRF | [8 empty] | 0
0
| 1 | 1 | 120 | [8 empty] | 1
| 1 | 2 | 110 | [8 empty] | 4
| 1 | 3 | 120 | [8 empty] | 16
| 1 | 4 | 110 | [8 empty] | 113.3333
| 1 | 5 | 110 | [8 empty] | 10
| 1 | 6 | 110
| Mean |  | 113.3333
| Median |  | 110
| Std Dev. |  | 5.164
| Rel.Std.Dev. |  | 4.556%
| PDM3 |  | 2.887e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 24b (Certified Value 2.46 wt.%) | [11 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 08 | 10 | 12 | 13 | 14 | wt.%
|  |  | BF*XRF | at 1000°C | BF*OES | BF*XRF | at 1000°C | at 1000°C | at 1000°C | BF*OES | BF*OES | BF*XRF | 2
3
| 1 | 1 | 2.33 | 2.29 | 2.76 | 2.59 | 2.29 | 2.28 | 2.36 | 2.5772 | 2.61 | 2.65 | 1
| 1 | 2 | 2.34 | 2.27 | 2.75 | 2.63 | 2.28 | 2.27 | 2.28 | 2.5938 | 2.6 | 2.66 | 0
| 1 | 3 | 2.4 | 2.27 | 2.72 | 2.62 | 2.27 | 2.3 | 2.31 | 2.5457 | 2.54 | 2.6 | 16
| 1 | 4 | 2.36 | 2.28 | 2.68 | 2.63 | 2.28 | 2.29 | 2.32 | 2.5339 | 2.7 | 2.62 | 2.4615
| 1 | 5 | 2.37 | 2.27 | 2.65 | 2.6 | 2.28 | 2.32 | 2.3 | 2.525 | 1.14 | 2.63 | 22
| 1 | 6 | 2.38 | 2.27 | 2.69 | 2.57 | 2.29 | 2.27 | 2.26 | 2.4688 | 2.64 | 2.61
| Mean |  | 2.3633 | 2.275 | 2.7083 | 2.6067 | 2.2817 | 2.2883 | 2.305 | 2.5407 | 2.3717 | 2.6283
| Median |  | 2.365 | 2.27 | 2.705 | 2.61 | 2.28 | 2.285 | 2.305 | 2.5398 | 2.605 | 2.625
| Std Dev. |  | 0.02582 | 0.0083666 | 0.042622 | 0.024221 | 0.0075277 | 0.019408 | 0.034496 | 0.043909 | 0.60565 | 0.023166
| Rel.Std.Dev. |  | 1.093% | 0.3678% | 1.574% | 0.9292% | 0.3299% | 0.8481% | 1.497% | 1.728% | 25.54% | 0.8814%
| PDM3 |  | -3.99% | -7.578% | 10.03% | 5.896% | -7.307% | -7.036% | -6.359% | 3.216% | -3.651% | 6.776%
| Z-Score (Absolute) |  | 0.03 | 0.7 | 2.58 | 1.81 | 0.65 | 0.6 | 0.47 | 1.31 | 0.03 | 1.97
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 24b (Indicative Value 2.17 ppm) | [14 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 05 | 06 | 07 | 12 | 13 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | [7 empty] | 2
2
| 1 | 1 | < 5 | < 2 | 0.33367 | 30 | 0.4105 | 2 | [7 empty] | 1
| 1 | 2 | < 5 | < 2 | 0.23663 | < 10 | 0.29044 | 1.8 | [7 empty] | 7
| 1 | 3 | < 5 | < 2 | 0.51491 | < 10 | 0.22972 | 1.9 | [7 empty] | 16
| 1 | 4 | < 5 | < 2 | 0.27173 | < 10 | 0.30431 | 2 | [7 empty] | 2.173
| 1 | 5 | 5 | < 2 | 0.22372 | < 10 | 0.43873 | 1.7 | [7 empty] | 13
| 1 | 6 | 5 | < 2 | 0.30137 | < 10 | 0.82573 | 1.9
| Mean |  | 5 |  | 0.31367 | 30 | 0.41657 | 1.8833
| Median |  | 5 |  | 0.28655 | 30 | 0.3574 | 1.9
| Std Dev. |  | 0 |  | 0.10661 |  | 0.21511 | 0.1169
| Rel.Std.Dev. |  | 0% |  | 33.99% |  | 51.64% | 6.207%
| PDM3 |  | 130.1% |  | -85.57% | 1281% | -80.83% | -13.33%
| Z-Score (Absolute) |  | 1.18 | 0.28 | 0.71 | 4.84 | 0.64 | 0.28
| Analytical results for Al in OREAS 24b (Certified Value 7.81 wt.%) | [14 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS | 2
3
| 1 | 1 | 7.621 | 7.86 | 7.817 | 7.39 | 7.3019 | 7.9 | 8.283 | 7.85 | 7.859 | 8.097 | 7.6015 | 7.732 | 7.912 | 1
| 1 | 2 | 7.568 | 8.12 | 7.796 | 8.01 | 7.3373 | 7.83 | 8.283 | 7.57 | 7.753 | 7.992 | 7.5454 | 8.039 | 7.833 | 0
| 1 | 3 | 7.648 | 7.88 | 7.833 | 8.01 | 7.3237 | 7.91 | 8.415 | 7.74 | 7.806 | 7.663 | 7.6496 | 8.055 | 7.812 | 16
| 1 | 4 | 7.648 | 7.95 | 7.865 | 9.21 | 7.3177 | 8.05 | 8.283 | 7.73 | 7.859 | 7.685 | 7.6147 | 7.912 | 7.806 | 7.8113
| 1 | 5 | 7.621 | 7.82 | 7.907 | 7.61 | 7.281 | 7.89 | 8.388 | 7.89 | 7.806 | 8.171 | 7.6428 | 7.933 | 7.812 | 24
| 1 | 6 | 7.568 | 7.85 | 7.87 | 8.47 | 7.2749 | 8.01 | 8.256 | 7.52 | 7.806 | 7.579 | 7.6089 | 7.87 | 7.817
| Mean |  | 7.6123 | 7.9133 | 7.848 | 8.1167 | 7.3061 | 7.9317 | 8.318 | 7.7167 | 7.8148 | 7.8645 | 7.6105 | 7.9235 | 7.832
| Median |  | 7.621 | 7.87 | 7.849 | 8.01 | 7.3098 | 7.905 | 8.283 | 7.735 | 7.806 | 7.8385 | 7.6118 | 7.9225 | 7.8145
| Std Dev. |  | 0.036401 | 0.11021 | 0.040358 | 0.65234 | 0.024663 | 0.082077 | 0.066073 | 0.14747 | 0.039897 | 0.25242 | 0.037179 | 0.11865 | 0.040254
| Rel.Std.Dev. |  | 0.4782% | 1.393% | 0.5143% | 8.037% | 0.3376% | 1.035% | 0.7943% | 1.911% | 0.5105% | 3.21% | 0.4885% | 1.497% | 0.514%
| PDM3 |  | -2.547% | 1.307% | 0.4703% | 3.91% | -6.467% | 1.541% | 6.487% | -1.211% | 0.04566% | 0.6815% | -2.57% | 1.437% | 0.2654%
| Z-Score (Absolute) |  | 1.89 | 0.53 | 0 | 2.16 | 4.35 | 0.67 | 3.78 | 1.05 | 0.26 | 0.13 | 1.91 | 0.61 | 0.13
| Analytical results for As in OREAS 24b (Indicative Value 9.74 ppm) | [14 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 13 |  |  |  |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | BF*MS |  |  |  |  | 2
2
| 1 | 1 | 100 | 20 | < 20 | 9.6069 | 19 | < 100 | < 70 | < 100 | 8 |  |  |  |  | 1
| 1 | 2 | < 100 | 20 | < 20 | 8.5465 | < 5 | < 100 | < 70 | < 100 | 10 |  |  |  |  | 3
| 1 | 3 | 100 | 10 | < 20 | 11.005 | < 5 | 100 | < 70 | < 100 | 10 |  |  |  |  | 16
| 1 | 4 | < 100 | 15 | < 20 | 8.2777 | 23 | < 100 | < 70 | < 100 | 10 |  |  |  |  | 9.7352
| 1 | 5 | < 100 | 5 | < 20 | 9.3616 | < 5 | < 100 | < 70 | < 100 | 9 |  |  |  |  | 14
| 1 | 6 | < 100 | 15 | < 20 | 9.0479 | < 5 | < 100 | < 70 | < 100 | 11
| Mean |  | 100 | 14.1667 |  | 9.3076 | 21 | 100 |  |  | 9.6667
| Median |  | 100 | 15 |  | 9.2048 | 21 | 100 |  |  | 10
| Std Dev. |  | 01: 0; 02: 5.8452; 06: 0.96756; 07: 2.8284; 13: 1.0328
| Rel.Std.Dev. |  | 01: 0%; 02: 41.26%; 06: 10.4%; 07: 13.47%; 13: 10.68%
| PDM3 |  | 927.2% | 45.52% |  | -4.393% | 115.7% | 927.2% |  |  | -0.7043%
| Z-Score (Absolute) |  | 6.44 | 0 | 0.51 | 0.6 | 0.67 | 5.41 | 2.55 | 4.39 | 0.55
| Analytical results for B in OREAS 24b (Indicative Value 69 ppm) | [14 empty] | Indicative
B | Round | Replicate | 01: Lab; 02: Lab |  |  |  |  | 1
| No | No | 01: 02; 02: 07 |  |  |  |  | ppm
|  |  | 01: PF*OES; 02: PF*MS |  |  |  |  | 0
0
| 1 | 1 | 01: 60; 02: 80 |  |  |  |  | 1
| 1 | 2 | 01: 60; 02: 80 |  |  |  |  | 9
| 1 | 3 | 01: 60; 02: 70 |  |  |  |  | 16
| 1 | 4 | 01: 80; 02: 80 |  |  |  |  | 69.1667
| 1 | 5 | 01: 40; 02: 80 |  |  |  |  | 15
| 1 | 6 | 01: 60; 02: 80
| Mean |  | 01: 60; 02: 78.3333
| Median |  | 01: 60; 02: 80
| Std Dev. |  | 01: 12.6491; 02: 4.0825
| Rel.Std.Dev. |  | 01: 21.08%; 02: 5.212%
| PDM3 |  | 01: -13.25%; 02: 13.25%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Ba in OREAS 24b (Certified Value 739 ppm) | [14 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*OES | BF*OES | BF*OES | BF*OES/MS |  |  | 0
0
| 1 | 1 | 786 | 730 | 725 | 672 | 739.69 | 756 | 744 | 711 | 754.3611 | 731 | 759 |  |  | 1
| 1 | 2 | 778 | 760 | 732 | 660 | 739 | 701 | 776 | 712 | 750.2049 | 736 | 732 |  |  | 0
| 1 | 3 | 789 | 740 | 729 | 678 | 733.92 | 714 | 785 | 667 | 751.1506 | 749 | 724 |  |  | 16
| 1 | 4 | 768 | 740 | 731 | 668 | 744.02 | 712 | 780 | 676 | 742.4926 | 743 | 728 |  |  | 738.9374
| 1 | 5 | 804 | 720 | 726 | 675 | 730.8 | 731 | 767 | 704 | 762.4038 | 752 | 738 |  |  | 25
| 1 | 6 | 794 | 720 | 717 | 673 | 741.06 | 705 | 800 | 668 | 751.343 | 749 | 729
| Mean |  | 786.5 | 735 | 726.6667 | 671 | 738.0817 | 719.8333 | 775.3333 | 689.6667 | 751.9927 | 743.3333 | 735
| Median |  | 787.5 | 735 | 727.5 | 672.5 | 739.345 | 713 | 778 | 690 | 751.2468 | 746 | 730.5
| Std Dev. |  | 12.518 | 15.1658 | 5.465 | 6.3246 | 4.8561 | 20.5077 | 18.8432 | 21.5839 | 6.4588 | 8.3106 | 12.6491
| Rel.Std.Dev. |  | 1.592% | 2.063% | 0.7521% | 0.9426% | 0.6579% | 2.849% | 2.43% | 3.13% | 0.8589% | 1.118% | 1.721%
| PDM3 |  | 6.437% | -0.5329% | -1.661% | -9.194% | -0.1158% | -2.585% | 4.925% | -6.668% | 1.767% | 0.5949% | -0.5329%
| Z-Score (Absolute) |  | 2.29 | 0 | 0.37 | 2.85 | 0.14 | 0.67 | 1.79 | 2.02 | 0.76 | 0.37 | 0
| Analytical results for Be in OREAS 24b (Certified Value 2.95 ppm) | [14 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 11 | 13 | 14 | [5 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*OES | BF*OES/MS | [5 empty] | 2
2
| 1 | 1 | 3 | 3.3 | 2 | 2.9064 | < 3 | 2.9 | 3 | 4 | [5 empty] | 1
| 1 | 2 | 3 | 3.2 | 3 | 2.8343 | < 3 | 2.8 | 3 | 3 | [5 empty] | 25
| 1 | 3 | 3 | 2.9 | 3 | 2.8548 | < 3 | 2.8 | 3 | 4 | [5 empty] | 16
| 1 | 4 | 3 | 3 | 3 | 2.8515 | < 3 | 3.2 | 3 | 2 | [5 empty] | 2.9461
| 1 | 5 | 3 | 3 | 3 | 2.8049 | < 3 | 2.1 | 3 | 2 | [5 empty] | 26
| 1 | 6 | 3 | 3.1 | 3 | 2.8853 | < 3 | 2.3 | 3 | 3
| Mean |  | 3 | 3.0833 | 2.8333 | 2.8562 |  | 2.6833 | 3 | 3
| Median |  | 3 | 3.05 | 3 | 2.8532 |  | 2.8 | 3 | 3
| Std Dev. |  | 0 | 0.1472 | 0.40825 | 0.036051 |  | 0.40702 | 0 | 0.89443
| Rel.Std.Dev. |  | 0% | 4.774% | 14.41% | 1.262% |  | 15.17% | 0% | 29.81%
| PDM3 |  | 1.829% | 4.657% | -3.828% | -3.052% |  | -8.92% | 1.829% | 1.829%
| Z-Score (Absolute) |  | 0.58 | 1.26 | 0.77 | 0.58 | 11.56 | 1.98 | 0.58 | 0.58
| Analytical results for Bi in OREAS 24b (Indicative Value 1.03 ppm) | [14 empty] | Indicative
Bi | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 13 |  |  | ppm
|  |  | 01: PF*MS; 02: BF*MS; 03: PF*MS; 05: PF*OES/MS; 06: PF*MS; 07: BF*MS |  |  | 2
2
| 1 | 1 | 01: < 2; 02: < 3; 03: 0.6; 05: 0.68; 06: < 2; 07: 1.5 |  |  | 1
| 1 | 2 | 01: < 2; 02: < 3; 03: 0.7; 05: 0.67; 06: < 2; 07: 1.6 |  |  | 10
| 1 | 3 | 01: < 2; 02: < 3; 03: 0.7; 05: 0.84; 06: < 2; 07: 1.6 |  |  | 16
| 1 | 4 | 01: < 2; 02: < 3; 03: 1; 05: 0.61; 06: < 2; 07: 1.8 |  |  | 1.0258
| 1 | 5 | 01: < 2; 02: < 3; 03: 0.7; 05: 0.59; 06: < 2; 07: 1.6 |  |  | 16
| 1 | 6 | 01: < 2; 02: < 3; 03: 1; 05: 0.67; 06: < 2; 07: 1.8
| Mean |  | 03: 0.78333; 05: 0.67667; 07: 1.65
| Median |  | 03: 0.7; 05: 0.67; 07: 1.6
| Std Dev. |  | 03: 0.17224; 05: 0.088015; 07: 0.12247
| Rel.Std.Dev. |  | 03: 21.99%; 05: 13.01%; 07: 7.423%
| PDM3 |  | 03: -23.64%; 05: -34.03%; 07: 60.85%
| Z-Score (Absolute) |  | 01: 0; 02: 1.25; 03: 0.54; 05: 0.81; 06: 0; 07: 1.62
| Analytical results for Ca in OREAS 24b (Certified Value 1.06 wt.%) | [14 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS |  | 2
3
| 1 | 1 | 1.015 | 1.029 | 1.03 | 1.062 | 1.08 | 1.108 | 1.022 | 1.072 | 1.115 | 1.0303 | 1.086 | 1.051 |  | 1
| 1 | 2 | 1.001 | 1.022 | 1.19 | 1.078 | 1.11 | 1.101 | 0.986 | 1.065 | 1.108 | 1.0357 | 1.093 | 1.043 |  | 0
| 1 | 3 | 1.001 | 1.015 | 1.2 | 1.082 | 1.05 | 1.136 | 1.013 | 1.093 | 1.051 | 1.0501 | 1.093 | 1.043 |  | 16
| 1 | 4 | 0.979 | 1.036 | 1.18 | 1.069 | 1.07 | 1.115 | 1.009 | 1.115 | 1.043 | 1.0509 | 1.086 | 1.036 |  | 1.0562
| 1 | 5 | 0.965 | 1.043 | 1.13 | 1.062 | 1.03 | 1.144 | 1.039 | 1.079 | 1.093 | 1.0402 | 1.079 | 1.051 |  | 27
| 1 | 6 | 0.972 | 1.036 | 1.18 | 1.074 | 1.03 | 1.151 | 0.971 | 1.079 | 1.043 | 1.0443 | 1.079 | 1.051
| Mean |  | 0.98883 | 1.0302 | 1.1517 | 1.0712 | 1.0617 | 1.1258 | 1.0067 | 1.0838 | 1.0755 | 1.0419 | 1.086 | 1.0458
| Median |  | 0.99 | 1.0325 | 1.18 | 1.0715 | 1.06 | 1.1255 | 1.011 | 1.079 | 1.072 | 1.0423 | 1.086 | 1.047
| Std Dev. |  | 0.019641 | 0.010304 | 0.064317 | 0.0083046 | 0.031252 | 0.020586 | 0.024598 | 0.017871 | 0.033572 | 0.0081111 | 0.006261 | 0.0062102
| Rel.Std.Dev. |  | 1.986% | 1% | 5.585% | 0.7753% | 2.944% | 1.828% | 2.444% | 1.649% | 3.122% | 0.7785% | 0.5765% | 0.5938%
| PDM3 |  | -6.375% | -2.461% | 9.042% | 1.421% | 0.5211% | 6.596% | -4.686% | 2.62% | 1.831% | -1.35% | 2.825% | -0.9781%
| Z-Score (Absolute) |  | 2.33 | 1.08 | 2.55 | 0.14 | 0.14 | 1.77 | 1.79 | 0.52 | 0.28 | 0.73 | 0.6 | 0.62
| Analytical results for Cd in OREAS 24b (Indicative Value < 2 ppm) | [14 empty] | Indicative
Cd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 06: Lab |  | 1
| No | No | 01: 02; 03: 05; 04: 06; 06: 07 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*MS |  | 2
2
| 1 | 1 | 01: < 10; 03: < 1; 04: 0.094758; 06: < 2 |  | 1
| 1 | 2 | 01: < 10; 03: < 1; 04: 0.15743; 06: < 2 |  | 11
| 1 | 3 | 01: < 10; 03: < 1; 04: 0.030815; 06: < 2 |  | 16
| 1 | 4 | 01: < 10; 03: < 1; 04: 0.025974; 06: < 2 |  | < 2
| 1 | 5 | 01: < 10; 03: < 1; 04: 0.050145; 06: < 2 |  | 17
| 1 | 6 | 01: < 10; 03: < 1; 04: 0.14873; 06: < 2
| Mean |  | 04: 0.084641
| Median |  | 04: 0.072451
| Std Dev. |  | 04: 0.058371
| Rel.Std.Dev. |  | 04: 68.96%
PDM3
| Z-Score (Absolute) |  | 01: 6.26; 03: 0.37; 04: 0.98; 06: 0.37
| Analytical results for Ce in OREAS 24b (Certified Value 86 ppm) | [14 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 0
1
| 1 | 1 | 89 | 85 | 87 | 86.8 | 84.7 | 103 | 86.6 | 84.3 | 86.65 | 82.7 | 85.4 |  |  | 1
| 1 | 2 | 88.5 | 86.5 | 87 | 84.2 | 82.43 | 85.3 | 88.4 | 82.4 | 86.58 | 83.4 | 85.4 |  |  | 27
| 1 | 3 | 89.6 | 86.5 | 87 | 85.1 | 83.72 | 89.1 | 87.8 | 85.8 | 87.23 | 82.4 | 83.2 |  |  | 16
| 1 | 4 | 87.9 | 89.5 | 88 | 85 | 82.61 | 85.5 | 90.8 | 84.4 | 86.01 | 81.3 | 85.9 |  |  | 85.9039
| 1 | 5 | 92.8 | 86.5 | 87 | 86.6 | 83.16 | 87.8 | 85.3 | 83.7 | 86.51 | 82.9 | 84.4 |  |  | 28
| 1 | 6 | 91.3 | 82 | 85 | 90.1 | 83.25 | 84.8 | 90.5 | 84.1 | 86.11 | 83.1 | 83.6
| Mean |  | 89.85 | 86 | 86.8333 | 86.3 | 83.3117 | 89.25 | 88.2333 | 84.1167 | 86.515 | 82.6333 | 84.65
| Median |  | 89.3 | 86.5 | 87 | 85.85 | 83.205 | 86.65 | 88.1 | 84.2 | 86.545 | 82.8 | 84.9
| Std Dev. |  | 1.8555 | 2.4495 | 0.98319 | 2.1128 | 0.82349 | 6.9377 | 2.1547 | 1.1017 | 0.43652 | 0.73666 | 1.0913
| Rel.Std.Dev. |  | 2.065% | 2.848% | 1.132% | 2.448% | 0.9884% | 7.773% | 2.442% | 1.31% | 0.5046% | 0.8915% | 1.289%
| PDM3 |  | 4.594% | 0.1118% | 1.082% | 0.4611% | -3.018% | 3.895% | 2.712% | -2.081% | 0.7113% | -3.807% | -1.46%
| Z-Score (Absolute) |  | 1.24 | 0.1 | 0.19 | 0 | 1.04 | 1.03 | 0.67 | 0.76 | 0.07 | 1.28 | 0.58
| Analytical results for Co in OREAS 24b (Certified Value 16.9 ppm) | [14 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 |  | ppm
|  |  | BF*MS | PF*OES | BF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  | 1
2
| 1 | 1 | 17.9 | < 20 | 18 | 17.26 | 16.4 | < 50 | 40 | 16.9 | 17.1 | 17.7976 | 17 | 15.6 |  | 1
| 1 | 2 | 17.8 | 20 | 17 | 16.13 | 16.2 | < 50 | 30 | 16.9 | 16.9 | 17.3033 | 16 | 15.2 |  | 5
| 1 | 3 | 17.9 | 20 | 17 | 17.3 | 16 | < 50 | 40 | 17.1 | 17 | 17.6168 | 16 | 16.9 |  | 16
| 1 | 4 | 17.8 | < 20 | 18 | 17.34 | 17 | < 50 | 40 | 16.9 | 17.3 | 17.4747 | 17 | 15.3 |  | 16.9403
| 1 | 5 | 18.2 | 20 | 17 | 17.16 | 16.4 | < 50 | 40 | 16.4 | 16.9 | 18.1904 | 17 | 15.3 |  | 29
| 1 | 6 | 17.9 | 20 | 18 | 17.49 | 17 | < 50 | 30 | 16.8 | 17.2 | 17.5508 | 16 | 14.5
| Mean |  | 17.9167 | 20 | 17.5 | 17.1133 | 16.5 |  | 36.6667 | 16.8333 | 17.0667 | 17.6556 | 16.5 | 15.4667
| Median |  | 17.9 | 20 | 17.5 | 17.28 | 16.4 |  | 40 | 16.9 | 17.05 | 17.5838 | 16.5 | 15.3
| Std Dev. |  | 0.1472 | 0 | 0.54772 | 0.49371 | 0.41473 |  | 5.164 | 0.23381 | 0.1633 | 0.30836 | 0.54772 | 0.79162
| Rel.Std.Dev. |  | 0.8216% | 0% | 3.13% | 2.885% | 2.514% |  | 14.08% | 1.389% | 0.9568% | 1.747% | 3.32% | 5.118%
| PDM3 |  | 5.764% | 18.06% | 3.304% | 1.022% | -2.599% |  | 116.4% | -0.6312% | 0.7462% | 4.223% | -2.599% | -8.699%
| Z-Score (Absolute) |  | 0.96 | 0.49 | 0.48 | 0.03 | 0.69 | 9.23 | 22.85 | 0.3 | 0.03 | 0.66 | 0.69 | 1.89
| Analytical results for Cr in OREAS 24b (Certified Value 142 ppm) | [14 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 |  | ppm
|  |  | BF*MS | PF*OES | BF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | BF*MS | BF*OES | BF*OES | BF*MS | BF*OES/MS |  | 0
0
| 1 | 1 | 180 | 200 | 125 | 134.5867 | 150 | 140 | 320 | 130 | 153 | 148.2251 | 130 | 137 |  | 1
| 1 | 2 | 170 | 200 | 123 | 134.1386 | 140 | 140 | 290 | 140 | 142 | 148.822 | 130 | 130 |  | 0
| 1 | 3 | 160 | 150 | 128 | 138.2306 | 150 | 210 | 190 | 150 | 132 | 145.6614 | 130 | 137 |  | 16
| 1 | 4 | 170 | 150 | 128 | 133.976 | 150 | 210 | 300 | 140 | 134 | 151.2731 | 130 | 137 |  | 141.7297
| 1 | 5 | 170 | 150 | 129 | 139.4118 | 140 | 210 | 380 | 130 | 143 | 151.1021 | 130 | 137 |  | 30
| 1 | 6 | 170 | 150 | 130 | 136.1155 | 140 | 210 | 480 | 130 | 131 | 146.8506 | 130 | 130
| Mean |  | 170 | 166.6667 | 127.1667 | 136.0765 | 145 | 186.6667 | 326.6667 | 136.6667 | 139.1667 | 148.6557 | 130 | 134.6667
| Median |  | 170 | 150 | 128 | 135.3511 | 145 | 210 | 310 | 135 | 138 | 148.5236 | 130 | 137
| Std Dev. |  | 6.3246 | 25.8199 | 2.6394 | 2.287 | 5.4772 | 36.1478 | 97.091 | 8.165 | 8.4715 | 2.2482 | 0 | 3.6148
| Rel.Std.Dev. |  | 3.72% | 15.49% | 2.076% | 1.681% | 3.777% | 19.36% | 29.72% | 5.974% | 6.087% | 1.512% | 0% | 2.684%
| PDM3 |  | 19.95% | 17.59% | -10.28% | -3.989% | 2.307% | 31.71% | 130.5% | -3.572% | -1.808% | 4.887% | -8.276% | -4.983%
| Z-Score (Absolute) |  | 1.92 | 1.69 | 1.03 | 0.41 | 0.2 | 2.78 | 12.7 | 0.37 | 0.2 | 0.45 | 0.83 | 0.52
| Analytical results for Cs in OREAS 24b (Certified Value 10.5 ppm) | [14 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 1
2
| 1 | 1 | 10.9 | 10 | 11 | 9.92 | 10.79 | 10.4 | 9.97 | 9.96 | 10.627 | 10.6 | 10.3 |  |  | 1
| 1 | 2 | 10.8 | 10 | 11 | 10.03 | 11.21 | 10.2 | 10.15 | 9.88 | 10.5634 | 10.8 | 10 |  |  | 0
| 1 | 3 | 11 | 10 | 11 | 10.39 | 11.13 | 9.9 | 10.6 | 10.01 | 10.7595 | 10.7 | 10.2 |  |  | 16
| 1 | 4 | 10.75 | 10 | 11 | 10.12 | 11.1 | 10.3 | 10.35 | 9.82 | 10.7107 | 10.9 | 9.9 |  |  | 10.4884
| 1 | 5 | 11.15 | 10 | 11 | 10.24 | 10.78 | 11 | 10.15 | 10.19 | 10.7473 | 10.7 | 10.5 |  |  | 31
| 1 | 6 | 11.05 | 10 | 11 | 10.76 | 11.17 | 10 | 10.45 | 10.08 | 10.7241 | 10.8 | 10
| Mean |  | 10.9417 | 10 | 11 | 10.2433 | 11.03 | 10.3 | 10.2783 | 9.99 | 10.6887 | 10.75 | 10.15
| Median |  | 10.95 | 10 | 11 | 10.18 | 11.115 | 10.25 | 10.25 | 9.985 | 10.7174 | 10.75 | 10.1
| Std Dev. |  | 0.15303 | 0 | 0 | 0.30124 | 0.19339 | 0.38987 | 0.23069 | 0.13446 | 0.07705 | 0.10488 | 0.22583
| Rel.Std.Dev. |  | 1.399% | 0% | 0% | 2.941% | 1.753% | 3.785% | 2.244% | 1.346% | 0.7209% | 0.9756% | 2.225%
| PDM3 |  | 4.322% | -4.656% | 4.878% | -2.336% | 5.164% | -1.796% | -2.003% | -4.752% | 1.91% | 2.495% | -3.226%
| Z-Score (Absolute) |  | 1.4 | 0.65 | 1.52 | 0.12 | 1.59 | 0 | 0.05 | 0.67 | 0.85 | 0.98 | 0.33
| Analytical results for Cu in OREAS 24b (Indicative Value 35.1 ppm) | [14 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | BF*OES | BF*MS | BF*MS |  |  |  | 1
1
| 1 | 1 | < 50 | < 50 | 35.8754 | 51 | < 100 | < 80 | < 50 | 38 | 33.6382 | 40 |  |  |  | 1
| 1 | 2 | < 50 | < 50 | 35.1588 | 49 | < 100 | < 80 | 70 | 38 | 33.6721 | 40 |  |  |  | 12
| 1 | 3 | < 50 | < 50 | 37.1296 | 46 | < 100 | 90 | < 50 | 35 | 32.8508 | 40 |  |  |  | 16
| 1 | 4 | < 50 | < 50 | 34.4056 | 53 | < 100 | < 80 | < 50 | 36 | 34.3753 | 40 |  |  |  | 35.1403
| 1 | 5 | 50 | < 50 | 35.2976 | 65 | < 100 | 110 | < 50 | 38 | 34.5908 | 40 |  |  |  | 18
| 1 | 6 | < 50 | < 50 | 36.3687 | 49 | < 100 | 90 | < 50 | 34 | 35.9295 | 40
| Mean |  | 50 |  | 35.706 | 52.1667 |  | 96.6667 | 70 | 36.5 | 34.1761 | 40
| Median |  | 50 |  | 35.5865 | 50 |  | 90 | 70 | 37 | 34.0237 | 40
| Std Dev. |  |  |  | 0.96455 | 6.7057 |  | 11.547 |  | 1.7607 | 1.0569 | 0
| Rel.Std.Dev. |  |  |  | 2.701% | 12.85% |  | 11.95% |  | 4.824% | 3.093% | 0%
| PDM3 |  | 42.29% |  | 1.61% | 48.45% |  | 175.1% | 99.2% | 3.869% | -2.744% | 13.83%
| Z-Score (Absolute) |  | 0.86 | 1.38 | 0.05 | 2 | 1.73 | 4.01 | 0.45 | 0.05 | 0.24 | 0.49
| Analytical results for Dy in OREAS 24b (Certified Value 5.83 ppm) | [14 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 5.99 | 5.5 | 5.8 | 6.069 | 6.1 | 5.69 | 5.73 | 6.1246 | 5.6 | 5.98 |  |  |  | 1
| 1 | 2 | 5.99 | 6 | 6 | 5.729 | 5.7 | 5.74 | 5.84 | 6.0524 | 5.6 | 5.73 |  |  |  | 17
| 1 | 3 | 6.01 | 5.5 | 6.4 | 6.036 | 5.9 | 5.89 | 6.1 | 5.8661 | 5.5 | 5.75 |  |  |  | 16
| 1 | 4 | 5.84 | 5.5 | 5.6 | 5.983 | 5.5 | 6.25 | 6.09 | 6.0618 | 5.6 | 5.24 |  |  |  | 5.8344
| 1 | 5 | 6.04 | 6 | 6.1 | 5.767 | 5.8 | 6.09 | 6.04 | 5.7451 | 5.5 | 5.81 |  |  |  | 32
| 1 | 6 | 5.86 | 5.5 | 6.3 | 5.893 | 5.6 | 5.68 | 5.8 | 6.1763 | 5.4 | 5.38
| Mean |  | 5.955 | 5.6667 | 6.0333 | 5.9128 | 5.7667 | 5.89 | 5.9333 | 6.0044 | 5.5333 | 5.6483
| Median |  | 5.99 | 5.5 | 6.05 | 5.938 | 5.75 | 5.815 | 5.94 | 6.0571 | 5.55 | 5.74
| Std Dev. |  | 0.083606 | 0.2582 | 0.30111 | 0.14138 | 0.21602 | 0.23503 | 0.16219 | 0.16493 | 0.08165 | 0.27996
| Rel.Std.Dev. |  | 1.404% | 4.556% | 4.991% | 2.391% | 3.746% | 3.99% | 2.734% | 2.747% | 1.476% | 4.956%
| PDM3 |  | 2.067% | -2.875% | 3.41% | 1.345% | -1.161% | 0.9532% | 1.696% | 2.914% | -5.16% | -3.189%
| Z-Score (Absolute) |  | 0.31 | 1.35 | 0.76 | 0.07 | 0.77 | 0.07 | 0.18 | 0.59 | 2.11 | 1.45
| Analytical results for Er in OREAS 24b (Certified Value 3.41 ppm) | [14 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 3.53 | 3.5 | 3.5 | 3.467 | 3.5 | 3.48 | 3.31 | 3.4574 | 3.3 | 3.38 |  |  |  | 1
| 1 | 2 | 3.53 | 3.5 | 3.2 | 3.284 | 3.4 | 3.42 | 3.41 | 3.4568 | 3.2 | 3.43 |  |  |  | 18
| 1 | 3 | 3.49 | 3.5 | 3.6 | 3.456 | 3.5 | 3.6 | 3.6 | 3.5327 | 3.2 | 3.39 |  |  |  | 16
| 1 | 4 | 3.46 | 3.5 | 3.5 | 3.431 | 3.3 | 3.32 | 3.39 | 3.4242 | 3.2 | 3.65 |  |  |  | 3.4128
| 1 | 5 | 3.57 | 3.5 | 3.5 | 3.168 | 3.4 | 3.15 | 3.39 | 3.5985 | 3.2 | 3.15 |  |  |  | 33
| 1 | 6 | 3.56 | 3 | 3.9 | 3.39 | 3.2 | 3.53 | 3.44 | 3.3859 | 3.2 | 3.39
| Mean |  | 3.5233 | 3.4167 | 3.5333 | 3.366 | 3.3833 | 3.4167 | 3.4233 | 3.4759 | 3.2167 | 3.3983
| Median |  | 3.53 | 3.5 | 3.5 | 3.4105 | 3.4 | 3.45 | 3.4 | 3.4571 | 3.2 | 3.39
| Std Dev. |  | 0.041793 | 0.20412 | 0.22509 | 0.1175 | 0.1169 | 0.16182 | 0.096678 | 0.077126 | 0.040825 | 0.15905
| Rel.Std.Dev. |  | 1.186% | 5.974% | 6.371% | 3.491% | 3.455% | 4.736% | 2.824% | 2.219% | 1.269% | 4.68%
| PDM3 |  | 3.238% | 0.1126% | 3.531% | -1.372% | -0.8641% | 0.1126% | 0.3079% | 1.848% | -5.748% | -0.4246%
| Z-Score (Absolute) |  | 1.71 | 0 | 1.87 | 0.81 | 0.54 | 0 | 0.11 | 0.95 | 3.21 | 0.29
| Analytical results for Eu in OREAS 24b (Certified Value 1.39 ppm) | [14 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.44 | 1.6 | 1.2 | 1.4808 | 1.7 | 1.37 | 1.3 | 1.3918 | 1.43 | 1.32 |  |  |  | 1
| 1 | 2 | 1.41 | 1.6 | 1.4 | 1.3762 | 1.4 | 1.46 | 1.3 | 1.3971 | 1.4 | 1.27 |  |  |  | 19
| 1 | 3 | 1.45 | 1.4 | 1.5 | 1.4364 | 1.5 | 1.47 | 1.33 | 1.4174 | 1.42 | 1.35 |  |  |  | 16
| 1 | 4 | 1.41 | 1.6 | 1.2 | 1.4246 | 1.4 | 1.44 | 1.26 | 1.3895 | 1.39 | 1.31 |  |  |  | 1.3938
| 1 | 5 | 1.45 | 1.4 | 1.3 | 1.5738 | 1.4 | 1.47 | 1.28 | 1.4902 | 1.42 | 1.28 |  |  |  | 34
| 1 | 6 | 1.46 | 1.4 | 1.9 | 1.4442 | 1.3 | 1.43 | 1.33 | 1.389 | 1.41 | 1.27
| Mean |  | 1.4367 | 1.5 | 1.4167 | 1.456 | 1.45 | 1.44 | 1.3 | 1.4125 | 1.4117 | 1.3
| Median |  | 1.445 | 1.5 | 1.35 | 1.4403 | 1.4 | 1.45 | 1.3 | 1.3944 | 1.415 | 1.295
| Std Dev. |  | 0.021602 | 0.10954 | 0.26394 | 0.066873 | 0.13784 | 0.037947 | 0.027568 | 0.039505 | 0.01472 | 0.032249
| Rel.Std.Dev. |  | 1.504% | 7.303% | 18.63% | 4.593% | 9.506% | 2.635% | 2.121% | 2.797% | 1.043% | 2.481%
| PDM3 |  | 3.078% | 7.622% | 1.643% | 4.465% | 4.034% | 3.317% | -6.728% | 1.342% | 1.284% | -6.728%
| Z-Score (Absolute) |  | 0.35 | 2.58 | 0.35 | 1.03 | 0.82 | 0.47 | 4.46 | 0.5 | 0.53 | 4.46
| Analytical results for Fe in OREAS 24b (Certified Value 4.45 wt.%) | [14 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS | 2
3
| 1 | 1 | 4.337 | 4.65 | 4.483 | 4.17 | 4.3751 | 4.51 | 4.7 | 4.42 | 4.43 | 4.497 | 4.367 | 4.476 | 4.4 | 1
| 1 | 2 | 4.365 | 4.68 | 4.497 | 4.39 | 4.378 | 4.51 | 4.637 | 4.29 | 4.41 | 4.49 | 4.3821 | 4.532 | 4.33 | 0
| 1 | 3 | 4.372 | 4.5 | 4.49 | 4.37 | 4.3683 | 4.55 | 4.777 | 4.33 | 4.45 | 4.469 | 4.4131 | 4.546 | 4.351 | 16
| 1 | 4 | 4.344 | 4.61 | 4.518 | 5.02 | 4.3909 | 4.53 | 4.686 | 4.31 | 4.41 | 4.469 | 4.4151 | 4.448 | 4.337 | 4.4546
| 1 | 5 | 4.323 | 4.54 | 4.49 | 4.15 | 4.3569 | 4.52 | 4.798 | 4.39 | 4.4 | 4.525 | 4.4 | 4.56 | 4.386 | 35
| 1 | 6 | 4.33 | 4.6 | 4.539 | 4.62 | 4.3589 | 4.54 | 4.686 | 4.29 | 4.38 | 4.49 | 4.3863 | 4.504 | 4.393
| Mean |  | 4.3452 | 4.5967 | 4.5028 | 4.4533 | 4.3714 | 4.5267 | 4.714 | 4.3383 | 4.4133 | 4.49 | 4.3939 | 4.511 | 4.3662
| Median |  | 4.3405 | 4.605 | 4.4935 | 4.38 | 4.3717 | 4.525 | 4.693 | 4.32 | 4.41 | 4.49 | 4.3931 | 4.518 | 4.3685
| Std Dev. |  | 0.019508 | 0.067132 | 0.021424 | 0.32623 | 0.012762 | 0.01633 | 0.061185 | 0.054559 | 0.024221 | 0.020765 | 0.018822 | 0.043151 | 0.030486
| Rel.Std.Dev. |  | 0.449% | 1.46% | 0.4758% | 7.326% | 0.2919% | 0.3607% | 1.298% | 1.258% | 0.5488% | 0.4625% | 0.4284% | 0.9566% | 0.6982%
| PDM3 |  | -2.457% | 3.189% | 1.082% | -0.02904% | -1.869% | 1.617% | 5.823% | -2.611% | -0.927% | 0.7941% | -1.363% | 1.265% | -1.986%
| Z-Score (Absolute) |  | 1 | 1.32 | 0.46 | 0 | 0.75 | 0.67 | 2.4 | 1.06 | 0.37 | 0.34 | 0.55 | 0.53 | 0.81
| Analytical results for Ga in OREAS 24b (Certified Value 20.2 ppm) | [14 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 1
2
| 1 | 1 | 20.5 | 19 | 21 | 21.1902 | 21.8 | 19.5 | 20.4 | 19.9569 | 19 | 18.9 |  |  |  | 1
| 1 | 2 | 20.6 | 19 | 19 | 20.7561 | 21.3 | 20.3 | 20.2 | 19.8153 | 19 | 19 |  |  |  | 0
| 1 | 3 | 20.8 | 19 | 21 | 20.7678 | 21.1 | 19.4 | 20.4 | 20.0723 | 19 | 19.4 |  |  |  | 16
| 1 | 4 | 20.5 | 19 | 20 | 21.0215 | 21.9 | 20 | 20.8 | 19.8118 | 20 | 19.5 |  |  |  | 20.1584
| 1 | 5 | 21.4 | 19 | 20 | 20.7067 | 22 | 20 | 21 | 20.2254 | 20 | 19.3 |  |  |  | 36
| 1 | 6 | 21.3 | 19 | 20 | 21.1686 | 21.7 | 19.3 | 20.4 | 20.0104 | 20 | 19.3
| Mean |  | 20.85 | 19 | 20.1667 | 20.9351 | 21.6333 | 19.75 | 20.5333 | 19.982 | 19.5 | 19.2333
| Median |  | 20.7 | 19 | 20 | 20.8946 | 21.75 | 19.75 | 20.4 | 19.9836 | 19.5 | 19.3
| Std Dev. |  | 0.40373 | 0 | 0.75277 | 0.21875 | 0.3559 | 0.40373 | 0.30111 | 0.1585 | 0.54772 | 0.23381
| Rel.Std.Dev. |  | 1.936% | 0% | 3.733% | 1.045% | 1.645% | 2.044% | 1.466% | 0.7932% | 2.809% | 1.216%
| PDM3 |  | 3.431% | -5.746% | 0.0411% | 3.853% | 7.317% | -2.026% | 1.86% | -0.875% | -3.266% | -4.589%
| Z-Score (Absolute) |  | 0.77 | 1.07 | 0.09 | 0.86 | 1.56 | 0.32 | 0.46 | 0.09 | 0.57 | 0.84
| Analytical results for Gd in OREAS 24b (Certified Value 6.27 ppm) | [14 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 6.34 | 8 | 6.3 | 6.8226 | 5.5 | 6.03 | 5.83 | 6.4527 | 6.1 | 6.52 |  |  |  | 1
| 1 | 2 | 6.41 | 8 | 6.3 | 6.8485 | 4.8 | 5.71 | 6.02 | 6.3336 | 6 | 6.01 |  |  |  | 20
| 1 | 3 | 6.33 | 8 | 6.8 | 6.8529 | 5 | 5.58 | 6.18 | 6.3402 | 6 | 5.9 |  |  |  | 16
| 1 | 4 | 6.29 | 8 | 6.9 | 7.048 | 4.7 | 5.86 | 5.97 | 6.2275 | 5.7 | 5.99 |  |  |  | 6.2658
| 1 | 5 | 6.38 | 8 | 6.7 | 6.7338 | 4.9 | 6.27 | 6.27 | 6.3993 | 6.1 | 5.62 |  |  |  | 37
| 1 | 6 | 6.44 | 8 | 6.6 | 7.0695 | 4.6 | 5.98 | 5.93 | 6.2801 | 5.9 | 6.09
| Mean |  | 6.365 | 8 | 6.6 | 6.8959 | 4.9167 | 5.905 | 6.0333 | 6.3389 | 5.9667 | 6.0217
| Median |  | 6.36 | 8 | 6.65 | 6.8507 | 4.85 | 5.92 | 5.995 | 6.3369 | 6 | 6
| Std Dev. |  | 0.055408 | 0 | 0.25298 | 0.1334 | 0.31885 | 0.24501 | 0.16354 | 0.08061 | 0.15055 | 0.29335
| Rel.Std.Dev. |  | 0.8705% | 0% | 3.833% | 1.935% | 6.485% | 4.149% | 2.711% | 1.272% | 2.523% | 4.872%
| PDM3 |  | 1.583% | 27.68% | 5.334% | 10.06% | -21.53% | -5.758% | -3.71% | 1.167% | -4.774% | -3.896%
| Z-Score (Absolute) |  | 0.48 | 4.89 | 1.12 | 1.91 | 3.42 | 0.76 | 0.41 | 0.41 | 0.59 | 0.44
| Analytical results for Ge in OREAS 24b (Indicative Value 1.64 ppm) | [14 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab |  |  |  | 1
| No | No | 01: 02; 02: 06; 05: 07; 06: 13 |  |  |  | ppm
|  |  | 01: PF*MS; 02: PF*OES/MS; 05: PF*MS; 06: BF*MS |  |  |  | 2
2
| 1 | 1 | 01: < 20; 02: 2.4133; 05: 0.8; 06: 2 |  |  |  | 1
| 1 | 2 | 01: < 20; 02: 2.3782; 05: < 0.7; 06: 2 |  |  |  | 32
| 1 | 3 | 01: < 20; 02: 2.172; 05: < 0.7; 06: 2 |  |  |  | 16
| 1 | 4 | 01: < 20; 02: 2.5335; 05: < 0.7; 06: 2 |  |  |  | 1.6399
| 1 | 5 | 01: < 20; 02: 2.5757; 05: < 0.7; 06: 2 |  |  |  | 13
| 1 | 6 | 01: < 20; 02: 2.3454; 05: < 0.7; 06: 3
| Mean |  | 02: 2.403; 05: 0.8; 06: 2.1667
| Median |  | 02: 2.3958; 05: 0.8; 06: 2
| Std Dev. |  | 02: 0.1444; 06: 0.40825
| Rel.Std.Dev. |  | 02: 6.009%; 06: 18.84%
| PDM3 |  | 02: 46.53%; 05: -51.22%; 06: 32.12%
| Z-Score (Absolute) |  | 01: 5.26; 02: 0.08; 05: 1.27; 06: 0.08
| Analytical results for Hf in OREAS 24b (Certified Value 6.15 ppm) | [14 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 2
3
| 1 | 1 | 6.3 | 6 | 5 | 6.4 | 6.3276 | < 10 | 6.3 | 7 | 6.7554 | 5.7 | 6.2 |  |  | 1
| 1 | 2 | 6.2 | 6.4 | 6 | 6.3 | 6.4683 | < 10 | 6.1 | 6.7 | 6.4807 | 5.6 | 5.5 |  |  | 33
| 1 | 3 | 6.1 | 6.2 | 5 | 6.4 | 6.4553 | < 10 | 6.4 | 7 | 6.4995 | 5.7 | 6.5 |  |  | 16
| 1 | 4 | 6 | 6.4 | 5 | 6.1 | 6.047 | < 10 | 6.2 | 6.7 | 6.427 | 5.6 | 5.6 |  |  | 6.1473
| 1 | 5 | 6.2 | 6 | 6 | 6.4 | 6.5043 | < 10 | 6.1 | 6.7 | 6.5744 | 5.6 | 5.6 |  |  | 38
| 1 | 6 | 6.2 | 6 | 6 | 6.3 | 6.5763 | < 10 | 6.3 | 6.9 | 6.2251 | 5.5 | 5.5
| Mean |  | 6.1667 | 6.1667 | 5.5 | 6.3167 | 6.3965 |  | 6.2333 | 6.8333 | 6.4937 | 5.6167 | 5.8167
| Median |  | 6.2 | 6.1 | 5.5 | 6.35 | 6.4618 |  | 6.25 | 6.8 | 6.4901 | 5.6 | 5.6
| Std Dev. |  | 0.10328 | 0.19664 | 0.54772 | 0.1169 | 0.18947 |  | 0.12111 | 0.15055 | 0.17424 | 0.075277 | 0.42622
| Rel.Std.Dev. |  | 1.675% | 3.189% | 9.959% | 1.851% | 2.962% |  | 1.943% | 2.203% | 2.683% | 1.34% | 7.328%
| PDM3 |  | 0.3144% | 0.3144% | -10.53% | 2.754% | 4.053% |  | 1.399% | 11.16% | 5.634% | -8.633% | -5.379%
| Z-Score (Absolute) |  | 0 | 0 | 1.37 | 0.31 | 0.47 | 2.41 | 0.14 | 1.37 | 0.67 | 1.13 | 0.72
| Analytical results for Ho in OREAS 24b (Certified Value 1.17 ppm) | [14 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.18 | 1.2 | 1.2 | 1.2742 | 1.2 | 1.12 | 1.13 | 1.2399 | 1.1 | 1.16 |  |  |  | 1
| 1 | 2 | 1.18 | 1.2 | 1.1 | 1.2733 | 1.2 | 1.18 | 1.16 | 1.2197 | 1.1 | 1.11 |  |  |  | 15
| 1 | 3 | 1.17 | 1.2 | 1.2 | 1.2674 | 1.2 | 1.24 | 1.17 | 1.2063 | 1.1 | 1.11 |  |  |  | 16
| 1 | 4 | 1.14 | 1.2 | 1.2 | 1.2677 | 1.1 | 1.29 | 1.13 | 1.2444 | 1.1 | 1.09 |  |  |  | 1.1715
| 1 | 5 | 1.16 | 1 | 1.2 | 1.1649 | 1.2 | 1.18 | 1.16 | 1.2276 | 1.1 | 1.17 |  |  |  | 39
| 1 | 6 | 1.18 | 1 | 1.3 | 1.2818 | 1.1 | 1.15 | 1.13 | 1.2127 | 1.1 | 1.13
| Mean |  | 1.1683 | 1.1333 | 1.2 | 1.2549 | 1.1667 | 1.1933 | 1.1467 | 1.2251 | 1.1 | 1.1283
| Median |  | 1.175 | 1.2 | 1.2 | 1.2705 | 1.2 | 1.18 | 1.145 | 1.2236 | 1.1 | 1.12
| Std Dev. |  | 0.016021 | 0.10328 | 0.063246 | 0.044394 | 0.05164 | 0.06186 | 0.018619 | 0.015081 | 2.4324e-16 | 0.031252
| Rel.Std.Dev. |  | 1.371% | 9.113% | 5.27% | 3.538% | 4.426% | 5.184% | 1.624% | 1.231% | 2.211e-14% | 2.77%
| PDM3 |  | -0.2671% | -3.255% | 2.436% | 7.121% | -0.4094% | 1.867% | -2.117% | 4.577% | -6.1% | -3.682%
| Z-Score (Absolute) |  | 0.02 | 0.69 | 0.66 | 1.77 | 0.02 | 0.52 | 0.42 | 1.16 | 1.37 | 0.79
| Analytical results for In in OREAS 24b (Indicative Value < 0.2 ppm) | [14 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 13 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | [7 empty] | 2
2
| 1 | 1 | < 0.2 | < 0.5 | < 0.1 | 0.069556 | < 0.2 | < 0.2 | [7 empty] | 1
| 1 | 2 | < 0.2 | < 0.5 | < 0.1 | 0.069307 | < 0.2 | < 0.2 | [7 empty] | 8
| 1 | 3 | < 0.2 | < 0.5 | < 0.1 | 0.059642 | < 0.2 | < 0.2 | [7 empty] | 16
| 1 | 4 | < 0.2 | < 0.5 | < 0.1 | 0.07992 | < 0.2 | < 0.2 | [7 empty] | < 0.2
| 1 | 5 | < 0.2 | < 0.5 | < 0.1 | 0.080039 | < 0.2 | < 0.2 | [7 empty] | 14
| 1 | 6 | < 0.2 | < 0.5 | < 0.1 | 0.089041 | < 0.2 | < 0.2
| Mean |  | 06: 0.074584
| Median |  | 06: 0.074738
| Std Dev. |  | 06: 0.010427
| Rel.Std.Dev. |  | 06: 13.98%
PDM3
| Z-Score (Absolute) |  | 0 | 7.96 | 2.65 | 1.35 | 0 | 0
| Analytical results for K in OREAS 24b (Certified Value 2.74 wt.%) | [14 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS |  |  | 2
3
| 1 | 1 | 2.7 | 2.764 | 2.695 | 2.7 | 2.889 | 2.78 | 2.6 | 2.789 | 2.7802 | 2.715 | 2.773 |  |  | 1
| 1 | 2 | 2.7 | 2.764 | 2.733 | 2.7 | 2.914 | 2.68 | 2.5 | 2.715 | 2.747 | 2.748 | 2.739 |  |  | 0
| 1 | 3 | 2.7 | 2.773 | 2.751 | 2.7 | 2.922 | 2.74 | 2.6 | 2.632 | 2.7951 | 2.798 | 2.756 |  |  | 16
| 1 | 4 | 2.7 | 2.789 | 2.735 | 2.7 | 2.889 | 2.75 | 2.6 | 2.681 | 2.7636 | 2.773 | 2.723 |  |  | 2.7401
| 1 | 5 | 2.7 | 2.798 | 2.731 | 2.7 | 2.93 | 2.81 | 2.6 | 2.889 | 2.7935 | 2.798 | 2.764 |  |  | 40
| 1 | 6 | 2.7 | 2.789 | 2.739 | 2.7 | 2.889 | 2.62 | 2.6 | 2.607 | 2.7769 | 2.798 | 2.773
| Mean |  | 2.7 | 2.7795 | 2.7307 | 2.7 | 2.9055 | 2.73 | 2.5833 | 2.7188 | 2.7761 | 2.7717 | 2.7547
| Median |  | 2.7 | 2.781 | 2.734 | 2.7 | 2.9015 | 2.745 | 2.6 | 2.698 | 2.7786 | 2.7855 | 2.76
| Std Dev. |  | 4.8648e-16 | 0.01446 | 0.018864 | 4.8648e-16 | 0.01877 | 0.069282 | 0.040825 | 0.10524 | 0.018354 | 0.034215 | 0.020047
| Rel.Std.Dev. |  | 1.802e-14% | 0.5202% | 0.6908% | 1.802e-14% | 0.646% | 2.538% | 1.58% | 3.871% | 0.6611% | 1.234% | 0.7277%
| PDM3 |  | -1.464% | 1.437% | -0.3451% | -1.464% | 6.035% | -0.3694% | -5.722% | -0.7769% | 1.311% | 1.151% | 0.5308%
| Z-Score (Absolute) |  | 0.67 | 1.08 | 0 | 0.67 | 3.84 | 0.01 | 3.24 | 0.26 | 1 | 0.89 | 0.53
| Analytical results for La in OREAS 24b (Certified Value 44 ppm) | [14 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*OES | BF*MS | BF*MS | BF*OES/MS |  |  | 1
2
| 1 | 1 | 45.6 | 45 | 45 | 45.9 | 43.29 | 48.6 | 43.3 | 44 | 44.15 | 40.6 | 44.9 |  |  | 1
| 1 | 2 | 45.4 | 46.5 | 45 | 45.2 | 42.13 | 42.6 | 44.6 | 43 | 43.63 | 41.2 | 43.1 |  |  | 25
| 1 | 3 | 45.9 | 46 | 44 | 43.9 | 43.16 | 44.3 | 44.8 | 43 | 44.48 | 41.8 | 43.4 |  |  | 16
| 1 | 4 | 44.6 | 45 | 45 | 44.9 | 42.79 | 43.5 | 45.3 | 43 | 43.59 | 40.4 | 43 |  |  | 44.0168
| 1 | 5 | 47.1 | 45 | 44 | 45.8 | 42.75 | 44.6 | 44.7 | 44 | 44.03 | 41.1 | 44.4 |  |  | 41
| 1 | 6 | 46.5 | 43 | 44 | 48 | 43.03 | 43.2 | 45.9 | 42 | 43.74 | 41 | 43.7
| Mean |  | 45.85 | 45.0833 | 44.5 | 45.6167 | 42.8583 | 44.4667 | 44.7667 | 43.1667 | 43.9367 | 41.0167 | 43.75
| Median |  | 45.75 | 45 | 44.5 | 45.5 | 42.91 | 43.9 | 44.75 | 43 | 43.885 | 41.05 | 43.55
| Std Dev. |  | 0.8735 | 1.2007 | 0.54772 | 1.3732 | 0.4132 | 2.1519 | 0.86641 | 0.75277 | 0.34708 | 0.4916 | 0.75565
| Rel.Std.Dev. |  | 1.905% | 2.663% | 1.231% | 3.01% | 0.9641% | 4.839% | 1.935% | 1.744% | 0.79% | 1.199% | 1.727%
| PDM3 |  | 4.165% | 2.423% | 1.098% | 3.635% | -2.632% | 1.022% | 1.704% | -1.931% | -0.1821% | -6.816% | -0.6062%
| Z-Score (Absolute) |  | 1.3 | 0.58 | 0.03 | 1.08 | 1.51 | 0 | 0.28 | 1.22 | 0.5 | 3.25 | 0.67
| Analytical results for Li in OREAS 24b (Certified Value 52 ppm) | [14 empty] | Certified
Li | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab |  |  | 1
| No | No | 01: 02; 02: 05; 03: 06; 05: 07; 06: 09 |  |  | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 05: PF*MS; 06: PF*OES |  |  | 0
1
| 1 | 1 | 01: 56; 02: 49; 03: 51.5897; 05: 53; 06: 50 |  |  | 1
| 1 | 2 | 01: 61; 02: 47; 03: 50.8713; 05: 52; 06: 50 |  |  | 0
| 1 | 3 | 01: 56; 02: 49; 03: 51.3946; 05: 51; 06: 50 |  |  | 16
| 1 | 4 | 01: 52; 02: 49; 03: 52.4525; 05: 52; 06: 50 |  |  | 51.8011
| 1 | 5 | 01: 56; 02: 49; 03: 51.8062; 05: 53; 06: 60 |  |  | 42
| 1 | 6 | 01: 59; 02: 48; 03: 51.9198; 05: 52; 06: 50
| Mean |  | 01: 56.6667; 02: 48.5; 03: 51.6724; 05: 52.1667; 06: 51.6667
| Median |  | 01: 56; 02: 49; 03: 51.6979; 05: 52; 06: 50
| Std Dev. |  | 01: 3.0768; 02: 0.83666; 03: 0.53169; 05: 0.75277; 06: 4.0825
| Rel.Std.Dev. |  | 01: 5.43%; 02: 1.725%; 03: 1.029%; 05: 1.443%; 06: 7.902%
| PDM3 |  | 01: 9.393%; 02: -6.373%; 03: -0.2486%; 05: 0.7056%; 06: -0.2596%
| Z-Score (Absolute) |  | 01: 6.81; 02: 4.33; 03: 0; 05: 0.67; 06: 0.01
| Analytical results for Lu in OREAS 24b (Certified Value 0.49 ppm) | [14 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 10 | 11 | 12 | 13 | 14 |  |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.52 | 0.4 | 0.54 | 0.55544 | 0.46 | 0.47 | 0.54077 | 0.52 | 0.51 |  |  |  |  | 1
| 1 | 2 | 0.5 | 0.6 | 0.49 | 0.50792 | 0.46 | 0.46 | 0.55974 | 0.52 | 0.47 |  |  |  |  | 17
| 1 | 3 | 0.49 | 0.4 | 0.5 | 0.53877 | 0.44 | 0.45 | 0.54545 | 0.52 | 0.47 |  |  |  |  | 16
| 1 | 4 | 0.48 | 0.4 | 0.49 | 0.56543 | 0.45 | 0.45 | 0.54401 | 0.52 | 0.46 |  |  |  |  | 0.48881
| 1 | 5 | 0.5 | 0.4 | 0.52 | 0.55641 | 0.48 | 0.46 | 0.54595 | 0.5 | 0.47 |  |  |  |  | 43
| 1 | 6 | 0.5 | 0.4 | 0.61 | 0.54012 | 0.46 | 0.48 | 0.51763 | 0.5 | 0.46
| Mean |  | 0.49833 | 0.43333 | 0.525 | 0.54402 | 0.45833 | 0.46167 | 0.54226 | 0.51333 | 0.47333
| Median |  | 0.5 | 0.4 | 0.51 | 0.54778 | 0.46 | 0.46 | 0.54473 | 0.52 | 0.47
| Std Dev. |  | 0.013292 | 0.08165 | 0.045935 | 0.020438 | 0.013292 | 0.01169 | 0.01372 | 0.010328 | 0.018619
| Rel.Std.Dev. |  | 2.667% | 18.84% | 8.749% | 3.757% | 2.9% | 2.532% | 2.53% | 2.012% | 3.934%
| PDM3 |  | 1.949% | -11.35% | 7.404% | 11.29% | -6.235% | -5.553% | 10.94% | 5.017% | -3.166%
| Z-Score (Absolute) |  | 0 | 1.2 | 0.49 | 0.84 | 0.74 | 0.67 | 0.81 | 0.28 | 0.46
| Analytical results for Mg in OREAS 24b (Certified Value 1.62 wt.%) | [14 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS | 2
3
| 1 | 1 | 1.586 | 1.71 | 1.64 | 1.46 | 1.5551 | 1.64 | 1.622 | 1.61 | 1.616 | 1.683 | 1.6164 | 1.598 | 1.664 | 1
| 1 | 2 | 1.58 | 1.71 | 1.634 | 1.62 | 1.5515 | 1.65 | 1.604 | 1.56 | 1.604 | 1.664 | 1.6012 | 1.61 | 1.646 | 0
| 1 | 3 | 1.592 | 1.67 | 1.658 | 1.61 | 1.5623 | 1.65 | 1.64 | 1.6 | 1.622 | 1.568 | 1.6002 | 1.616 | 1.652 | 16
| 1 | 4 | 1.592 | 1.71 | 1.67 | 1.57 | 1.5736 | 1.64 | 1.616 | 1.58 | 1.604 | 1.574 | 1.5867 | 1.592 | 1.646 | 1.6185
| 1 | 5 | 1.58 | 1.7 | 1.695 | 1.55 | 1.5538 | 1.65 | 1.646 | 1.58 | 1.61 | 1.646 | 1.6258 | 1.598 | 1.652 | 44
| 1 | 6 | 1.574 | 1.71 | 1.689 | 1.6 | 1.5713 | 1.66 | 1.616 | 1.55 | 1.598 | 1.562 | 1.6062 | 1.598 | 1.658
| Mean |  | 1.584 | 1.7017 | 1.6643 | 1.5683 | 1.5613 | 1.6483 | 1.624 | 1.58 | 1.609 | 1.6162 | 1.6061 | 1.602 | 1.653
| Median |  | 1.583 | 1.71 | 1.664 | 1.585 | 1.5587 | 1.65 | 1.619 | 1.58 | 1.607 | 1.61 | 1.6037 | 1.598 | 1.652
| Std Dev. |  | 0.0072664 | 0.016021 | 0.025033 | 0.059133 | 0.0094232 | 0.0075277 | 0.01595 | 0.022804 | 0.0088318 | 0.054179 | 0.013636 | 0.0090333 | 0.0070143
| Rel.Std.Dev. |  | 0.4587% | 0.9415% | 1.504% | 3.77% | 0.6036% | 0.4567% | 0.9821% | 1.443% | 0.5489% | 3.352% | 0.849% | 0.5639% | 0.4243%
| PDM3 |  | -2.134% | 5.136% | 2.829% | -3.102% | -3.538% | 1.841% | 0.3372% | -2.381% | -0.5896% | -0.1468% | -0.7694% | -1.022% | 2.129%
| Z-Score (Absolute) |  | 0.58 | 2.14 | 1.28 | 0.94 | 1.11 | 0.91 | 0.35 | 0.67 | 0 | 0.16 | 0.07 | 0.16 | 1.02
| Analytical results for Mn in OREAS 24b (Certified Value 0.046 wt.%) | [14 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 |  |  | wt.%
|  |  | PF*OES | BF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS |  |  | 3
3
| 1 | 1 | 0.046 | 0.046 | 0.045961 | 0.0487 | 0.046 | 0.047 | 0.046 | 0.046 | 0.045723 | 0.0457 | 0.046 |  |  | 1
| 1 | 2 | 0.046 | 0.046 | 0.04513 | 0.0474 | 0.046 | 0.044 | 0.046 | 0.046 | 0.045475 | 0.0465 | 0.046 |  |  | 0
| 1 | 3 | 0.046 | 0.046 | 0.044998 | 0.0475 | 0.046 | 0.045 | 0.046 | 0.039 | 0.045328 | 0.0465 | 0.046 |  |  | 16
| 1 | 4 | 0.046 | 0.046 | 0.045399 | 0.0494 | 0.046 | 0.045 | 0.046 | 0.039 | 0.04591 | 0.0457 | 0.046 |  |  | 0.046175
| 1 | 5 | 0.046 | 0.046 | 0.044868 | 0.0479 | 0.046 | 0.045 | 0.046 | 0.046 | 0.045753 | 0.0465 | 0.046 |  |  | 45
| 1 | 6 | 0.046 | 0.046 | 0.045686 | 0.0483 | 0.046 | 0.044 | 0.046 | 0.039 | 0.047503 | 0.0457 | 0.046
| Mean |  | 0.046 | 0.046 | 0.04534 | 0.0482 | 0.046 | 0.045 | 0.046 | 0.0425 | 0.045949 | 0.0461 | 0.046
| Median |  | 0.046 | 0.046 | 0.045264 | 0.0481 | 0.046 | 0.045 | 0.046 | 0.0425 | 0.045738 | 0.0461 | 0.046
| Std Dev. |  | 7.6012e-18 | 7.6012e-18 | 0.00042247 | 0.0007642 | 7.6012e-18 | 0.0010954 | 7.6012e-18 | 0.0038341 | 0.00078957 | 0.00043818 | 7.6012e-18
| Rel.Std.Dev. |  | 1.652e-14% | 1.652e-14% | 0.9318% | 1.585% | 1.652e-14% | 2.434% | 1.652e-14% | 9.021% | 1.718% | 0.9505% | 1.652e-14%
| PDM3 |  | -0.3793% | -0.3793% | -1.808% | 4.385% | -0.3793% | -2.545% | -0.3793% | -7.959% | -0.4907% | -0.1627% | -0.3793%
| Z-Score (Absolute) |  | 0 | 0 | 1.96 | 3.02 | 0 | 2.56 | 0 | 6.74 | 0.91 | 0.67 | 0
| Analytical results for Mo in OREAS 24b (Indicative Value 4.91 ppm) | [14 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 06 | 07 | 10 | 11 | 12 | 13 |  |  |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS |  |  |  |  | 2
2
| 1 | 1 | 5 | 5 | 5 | 7.1583 | 5 | 6 | 5 | 4.2596 | 4 |  |  |  |  | 1
| 1 | 2 | 4 | 5 | 5 | 6 | < 1 | 6 | 5 | 4.2574 | 4 |  |  |  |  | 9
| 1 | 3 | 4 | 5 | 5 | 5.6113 | 3 | 8 | 5 | 4.2942 | 4 |  |  |  |  | 16
| 1 | 4 | 4 | 5 | 5 | 8.4196 | < 1 | 7 | 4 | 4.1676 | 4 |  |  |  |  | 4.9107
| 1 | 5 | 5 | 5 | 5 | 9 | 2 | 6 | 5 | 4.4737 | 4 |  |  |  |  | 15
| 1 | 6 | 5 | 5 | 5 | 6.2534 | 4 | 6 | 5 | 4.3921 | 4
| Mean |  | 4.5 | 5 | 5 | 7.0738 | 3.5 | 6.5 | 4.8333 | 4.3074 | 4
| Median |  | 4.5 | 5 | 5 | 6.7058 | 3.5 | 6 | 5 | 4.2769 | 4
| Std Dev. |  | 0.54772 | 0 | 0 | 1.3779 | 1.291 | 0.83666 | 0.40825 | 0.10895 | 0
| Rel.Std.Dev. |  | 12.17% | 0% | 0% | 19.48% | 36.89% | 12.87% | 8.447% | 2.529% | 0%
| PDM3 |  | -8.363% | 1.819% | 1.819% | 44.05% | -28.73% | 32.37% | -1.575% | -12.28% | -18.54%
| Z-Score (Absolute) |  | 0.43 | 0.21 | 0.21 | 2.87 | 2.99 | 2.14 | 0 | 0.67 | 1.07
| Analytical results for Na in OREAS 24b (Certified Value 0.824 wt.%) | [14 empty] | Certified
Na | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 06: Lab; 07: Lab |  |  |  |  | 1
| No | No | 01: 03; 02: 11; 03: 12; 06: 13; 07: 14 |  |  |  |  | wt.%
|  |  | 01: BF*OES; 02: BF*OES; 03: BF*OES; 06: BF*OES; 07: BF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 01: 0.838; 02: 0.861; 03: 0.80396; 06: 0.794; 07: 0.809 |  |  |  |  | 1
| 1 | 2 | 01: 0.838; 02: 0.846; 03: 0.80297; 06: 0.801; 07: 0.801 |  |  |  |  | 0
| 1 | 3 | 01: 0.838; 02: 0.853; 03: 0.81391; 06: 0.816; 07: 0.801 |  |  |  |  | 16
| 1 | 4 | 01: 0.853; 02: 0.853; 03: 0.80834; 06: 0.809; 07: 0.794 |  |  |  |  | 0.82354
| 1 | 5 | 01: 0.853; 02: 0.861; 03: 0.81077; 06: 0.831; 07: 0.801 |  |  |  |  | 46
| 1 | 6 | 01: 0.846; 02: 0.838; 03: 0.7999; 06: 0.823; 07: 0.809
| Mean |  | 01: 0.84433; 02: 0.852; 03: 0.80664; 06: 0.81233; 07: 0.8025
| Median |  | 01: 0.842; 02: 0.853; 03: 0.80615; 06: 0.8125; 07: 0.801
| Std Dev. |  | 01: 0.0073937; 02: 0.0088994; 03: 0.005272; 06: 0.013794; 07: 0.0057184
| Rel.Std.Dev. |  | 01: 0.8757%; 02: 1.045%; 03: 0.6536%; 06: 1.698%; 07: 0.7126%
| PDM3 |  | 01: 2.524%; 02: 3.455%; 03: -2.052%; 06: -1.361%; 07: -2.555%
| Z-Score (Absolute) |  | 01: 2.19; 02: 2.7; 03: 0.39; 06: 0; 07: 0.67
| Analytical results for Nb in OREAS 24b (Certified Value 16 ppm) | [14 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 1
2
| 1 | 1 | 15.8 | 15 | 14 | 23 | 16.8337 | 17.3 | 16 | 16.5 | 16.0947 | 16 | 14.9 |  |  | 1
| 1 | 2 | 15.8 | 15 | 14 | 21 | 16.604 | 16 | 16.4 | 15.9 | 15.7647 | 16 | 15.5 |  |  | 27
| 1 | 3 | 15.7 | 15 | 13 | 28 | 16.8201 | 15.6 | 17.1 | 16.3 | 16.1183 | 16 | 14.3 |  |  | 16
| 1 | 4 | 15.6 | 15 | 14 | 20 | 17.044 | 16.4 | 16.6 | 16 | 15.9316 | 16 | 15.7 |  |  | 15.9957
| 1 | 5 | 16.3 | 15 | 14 | 21 | 16.6905 | 16.5 | 16.4 | 16 | 15.9387 | 16 | 16 |  |  | 47
| 1 | 6 | 16.2 | 10 | 13 | 20 | 16.8787 | 16.8 | 16.9 | 16 | 15.6515 | 15 | 14.9
| Mean |  | 15.9 | 14.1667 | 13.6667 | 22.1667 | 16.8118 | 16.4333 | 16.5667 | 16.1167 | 15.9166 | 15.8333 | 15.2167
| Median |  | 15.8 | 15 | 14 | 21 | 16.8269 | 16.45 | 16.5 | 16 | 15.9352 | 16 | 15.2
| Std Dev. |  | 0.28284 | 2.0412 | 0.5164 | 3.0605 | 0.15286 | 0.59554 | 0.39328 | 0.23166 | 0.18247 | 0.40825 | 0.62743
| Rel.Std.Dev. |  | 1.779% | 14.41% | 3.779% | 13.81% | 0.9093% | 3.624% | 2.374% | 1.437% | 1.146% | 2.578% | 4.123%
| PDM3 |  | -0.5986% | -11.43% | -14.56% | 38.58% | 5.102% | 2.736% | 3.569% | 0.756% | -0.4949% | -1.015% | -4.871%
| Z-Score (Absolute) |  | 0.02 | 1.82 | 2.33 | 6.48 | 0.93 | 0.54 | 0.67 | 0.21 | 0 | 0.09 | 0.73
| Analytical results for Nd in OREAS 24b (Certified Value 38.7 ppm) | [14 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 1
2
| 1 | 1 | 37.8 | 39.5 | 38.7 | 39.4758 | 44 | 37.6 | 38.2 | 38.01 | 39.3 | 34.9 |  |  |  | 1
| 1 | 2 | 38.7 | 40 | 38 | 38.4238 | 32.9 | 38.9 | 37.6 | 37.71 | 38.5 | 35.9 |  |  |  | 18
| 1 | 3 | 38.1 | 40.5 | 39.3 | 39.2863 | 34.6 | 38.9 | 39.3 | 38.03 | 38.7 | 33.9 |  |  |  | 16
| 1 | 4 | 37.6 | 39.5 | 38.8 | 39.018 | 33 | 40 | 38.2 | 37.11 | 38.3 | 34.4 |  |  |  | 38.6759
| 1 | 5 | 39.5 | 38.5 | 38.9 | 38.2102 | 34.4 | 38.1 | 38.1 | 37.65 | 39.1 | 33.9 |  |  |  | 48
| 1 | 6 | 38.8 | 35.5 | 40 | 39.3415 | 32.7 | 39.6 | 37.7 | 37.48 | 38.8 | 33.1
| Mean |  | 38.4167 | 38.9167 | 38.95 | 38.9593 | 35.2667 | 38.85 | 38.1833 | 37.665 | 38.7833 | 34.35
| Median |  | 38.4 | 39.5 | 38.85 | 39.1521 | 33.7 | 38.9 | 38.15 | 37.68 | 38.75 | 34.15
| Std Dev. |  | 0.71391 | 1.8005 | 0.66558 | 0.52367 | 4.3542 | 0.8961 | 0.6047 | 0.34547 | 0.37103 | 0.96695
| Rel.Std.Dev. |  | 1.858% | 4.626% | 1.709% | 1.344% | 12.35% | 2.307% | 1.584% | 0.9172% | 0.9567% | 2.815%
| PDM3 |  | -0.6704% | 0.6224% | 0.7086% | 0.7325% | -8.815% | 0.45% | -1.274% | -2.614% | 0.2777% | -11.19%
| Z-Score (Absolute) |  | 0.35 | 0.6 | 0.67 | 0.68 | 6.34 | 0.48 | 0.79 | 1.78 | 0.35 | 8.08
| Analytical results for Ni in OREAS 24b (Certified Value 61 ppm) | [14 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | BF*OES | BF*MS | BF*MS | BF*OES/MS |  |  |  | 0
0
| 1 | 1 | 60 | 50 | 65.997 | 60 | 70 | 60 | 63 | 64.0281 | 70 | 54 |  |  |  | 1
| 1 | 2 | 60 | 50 | 66.0139 | 60 | 60 | 60 | 54 | 67.4982 | 70 | 59 |  |  |  | 3
| 1 | 3 | 80 | 100 | 64.4791 | 70 | 60 | 50 | 54 | 65.7507 | 70 | 56 |  |  |  | 16
| 1 | 4 | 70 | 50 | 64.6933 | 60 | 70 | 60 | 54 | 62.9616 | 70 | 57 |  |  |  | 61.4693
| 1 | 5 | 60 | < 50 | 64.4147 | 60 | 70 | 60 | 57 | 63.5755 | 70 | 54 |  |  |  | 49
| 1 | 6 | 80 | 50 | 63.1918 | 70 | 60 | 70 | 53 | 64.1525 | 70 | 57
| Mean |  | 68.3333 | 60 | 64.7983 | 63.3333 | 65 | 60 | 55.8333 | 64.6611 | 70 | 56.1667
| Median |  | 65 | 50 | 64.5862 | 60 | 65 | 60 | 54 | 64.0903 | 70 | 56.5
| Std Dev. |  | 9.8319 | 22.3607 | 1.0729 | 5.164 | 5.4772 | 6.3246 | 3.7639 | 1.6711 | 0 | 1.9408
| Rel.Std.Dev. |  | 14.39% | 37.27% | 1.656% | 8.154% | 8.427% | 10.54% | 6.741% | 2.584% | 0% | 3.455%
| PDM3 |  | 11.17% | -2.39% | 5.416% | 3.033% | 5.744% | -2.39% | -9.169% | 5.193% | 13.88% | -8.626%
| Z-Score (Absolute) |  | 0.7 | 1.59 | 0.13 | 0.11 | 0.16 | 0.65 | 1.32 | 0.11 | 0.97 | 1.27
| Analytical results for P in OREAS 24b (Indicative Value 0.073 wt.%) | [14 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 06 | 07 | 11 | 12 | 13 | 14 | [6 empty] | wt.%
|  |  | BF*OES | PF*OES/MS | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.061 | 0.0749 | 0.067 | 0.074 | 0.0624 | 0.074 | 0.074 | [6 empty] | 1
| 1 | 2 | 0.057 | 0.073146 | 0.067 | 0.07 | 0.0628 | 0.074 | 0.074 | [6 empty] | 10
| 1 | 3 | 0.061 | 0.072526 | 0.067 | 0.074 | 0.0624 | 0.07 | 0.074 | [6 empty] | 16
| 1 | 4 | 0.057 | 0.073228 | 0.067 | 0.074 | 0.0641 | 0.07 | 0.07 | [6 empty] | 0.072933
| 1 | 5 | 0.061 | 0.073922 | 0.066 | 0.074 | 0.0637 | 0.074 | 0.07 | [6 empty] | 16
| 1 | 6 | 0.057 | 0.067758 | 0.068 | 0.07 | 0.0646 | 0.074 | 0.074
| Mean |  | 0.059 | 0.07258 | 0.067 | 0.072667 | 0.063333 | 0.072667 | 0.072667
| Median |  | 0.059 | 0.073187 | 0.067 | 0.074 | 0.06325 | 0.074 | 0.074
| Std Dev. |  | 0.0021909 | 0.002497 | 0.00063246 | 0.0020656 | 0.0009331 | 0.0020656 | 0.0020656
| Rel.Std.Dev. |  | 3.713% | 3.44% | 0.944% | 2.843% | 1.473% | 2.843% | 2.843%
| PDM3 |  | -19.1% | -0.4842% | -8.135% | -0.3653% | -13.16% | -0.3653% | -0.3653%
| Z-Score (Absolute) |  | 61.67 | 0 | 25.17 | 0.67 | 41.65 | 0.67 | 0.67
| Analytical results for Pb in OREAS 24b (Indicative Value 22.9 ppm) | [14 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 |  |  |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | BF*MS | BF*MS |  |  |  | 1
1
| 1 | 1 | < 100 | < 20 | 28 | 22.2867 | 25.8 | < 100 | < 300 | < 100 | 24 | 20 |  |  |  | 1
| 1 | 2 | < 100 | < 20 | 25 | 21.4794 | 30.7 | < 100 | < 300 | < 100 | 23 | 21 |  |  |  | 11
| 1 | 3 | < 100 | < 20 | 27 | 23.2977 | 22.3 | < 100 | < 300 | < 100 | 25 | 22 |  |  |  | 16
| 1 | 4 | < 100 | 20 | 28 | 21.5928 | 29.5 | < 100 | < 300 | < 100 | 24 | 21 |  |  |  | 22.894
| 1 | 5 | < 100 | 20 | 27 | 21.4497 | 36.3 | 100 | < 300 | < 100 | 25 | 21 |  |  |  | 17
| 1 | 6 | < 100 | 20 | 34 | 22.0787 | 23.4 | < 100 | < 300 | < 100 | 24 | 22
| Mean |  |  | 20 | 28.1667 | 22.0308 | 28 | 100 |  |  | 24.1667 | 21.1667
| Median |  |  | 20 | 27.5 | 21.8357 | 27.65 | 100 |  |  | 24 | 21
| Std Dev. |  |  | 0 | 3.0605 | 0.708 | 5.233 |  |  |  | 0.75277 | 0.75277
| Rel.Std.Dev. |  |  | 0% | 10.87% | 3.214% | 18.69% |  |  |  | 3.115% | 3.556%
| PDM3 |  |  | -12.64% | 23.03% | -3.77% | 22.3% | 336.8% |  |  | 5.559% | -7.545%
| Z-Score (Absolute) |  | 1.48 | 0.88 | 0.01 | 0.41 | 0.01 | 2.04 | 8.22 | 1.48 | 0.26 | 0.47
| Analytical results for Pr in OREAS 24b (Certified Value 10.2 ppm) | [14 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 1
2
| 1 | 1 | 10.05 | 10.8 | 10.71 | 10.6935 | 12.4 | 10.1 | 10.01 | 10.2774 | 9.94 | 9.89 |  |  |  | 1
| 1 | 2 | 9.98 | 10.6 | 10.55 | 10.4931 | 10.1 | 10.05 | 9.8 | 10.2077 | 9.93 | 9.68 |  |  |  | 19
| 1 | 3 | 10.05 | 10.8 | 10.54 | 10.6362 | 10.5 | 10.4 | 10.28 | 10.1554 | 9.75 | 9.4 |  |  |  | 16
| 1 | 4 | 9.89 | 10.6 | 10.52 | 10.5624 | 10 | 10.35 | 10.15 | 10.2331 | 9.79 | 9.42 |  |  |  | 10.1851
| 1 | 5 | 10.3 | 10.4 | 10.87 | 10.4822 | 10.4 | 9.97 | 10.11 | 10.1783 | 9.82 | 9.43 |  |  |  | 50
| 1 | 6 | 10.15 | 10 | 11.59 | 10.5489 | 10 | 10.5 | 9.97 | 10.0892 | 9.9 | 9.26
| Mean |  | 10.07 | 10.5333 | 10.7967 | 10.5694 | 10.5667 | 10.2283 | 10.0533 | 10.1902 | 9.855 | 9.5133
| Median |  | 10.05 | 10.6 | 10.63 | 10.5557 | 10.25 | 10.225 | 10.06 | 10.193 | 9.86 | 9.425
| Std Dev. |  | 0.14184 | 0.30111 | 0.41123 | 0.082192 | 0.92232 | 0.21591 | 0.16549 | 0.065278 | 0.079183 | 0.22906
| Rel.Std.Dev. |  | 1.409% | 2.859% | 3.809% | 0.7776% | 8.729% | 2.111% | 1.646% | 0.6406% | 0.8035% | 2.408%
| PDM3 |  | -1.13% | 3.419% | 6.005% | 3.773% | 3.746% | 0.4245% | -1.294% | 0.05023% | -3.241% | -6.596%
| Z-Score (Absolute) |  | 0.28 | 0.64 | 1.17 | 0.72 | 0.71 | 0.04 | 0.31 | 0.04 | 0.7 | 1.38
| Analytical results for Rb in OREAS 24b (Certified Value 161 ppm) | [14 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 0
0
| 1 | 1 | 177.5 | 171 | 161.8 | 162.7 | 162.95 | 163 | 169 | 173.6 | 158.47 | 159 | 157.8 |  |  | 1
| 1 | 2 | 179 | 190 | 160.8 | 159.4 | 158.37 | 160 | 167 | 173.1 | 158.99 | 161 | 156.3 |  |  | 0
| 1 | 3 | 178 | 189 | 162.8 | 163.1 | 159.67 | 156 | 169 | 176.3 | 161.62 | 160 | 160.5 |  |  | 16
| 1 | 4 | 177.5 | 181 | 163.5 | 162 | 161.52 | 163 | 167 | 178 | 157.87 | 159 | 158.6 |  |  | 161.085
| 1 | 5 | 183.5 | 179 | 161.6 | 162.9 | 159.3 | 162 | 167.5 | 176.2 | 159.85 | 159 | 155.1 |  |  | 51
| 1 | 6 | 181.5 | 179 | 159.4 | 163.5 | 161.56 | 159 | 168 | 174.3 | 158.31 | 158 | 155.3
| Mean |  | 179.5 | 181.5 | 161.65 | 162.2667 | 160.5617 | 160.5 | 167.9167 | 175.25 | 159.185 | 159.3333 | 157.2667
| Median |  | 178.5 | 180 | 161.7 | 162.8 | 160.595 | 161 | 167.75 | 175.25 | 158.73 | 159 | 157.05
| Std Dev. |  | 2.4698 | 7.0922 | 1.4529 | 1.4895 | 1.721 | 2.7386 | 0.91742 | 1.8876 | 1.3723 | 1.0328 | 2.0983
| Rel.Std.Dev. |  | 1.376% | 3.908% | 0.8988% | 0.9179% | 1.072% | 1.706% | 0.5464% | 1.077% | 0.8621% | 0.6482% | 1.334%
| PDM3 |  | 11.43% | 12.67% | 0.3507% | 0.7336% | -0.3249% | -0.3632% | 4.241% | 8.793% | -1.18% | -1.087% | -2.37%
| Z-Score (Absolute) |  | 4.88 | 5.43 | 0 | 0.17 | 0.3 | 0.31 | 1.71 | 3.72 | 0.67 | 0.63 | 1.2
| Analytical results for Re in OREAS 24b (Indicative Value < 0.1 ppm) | [14 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 05 |  |  | ppm
|  |  | 01: PF*MS; 02: BF*MS; 03: PF*MS |  |  | 2
2
| 1 | 1 | 01: < 1; 02: < 0.1; 03: < 0.1 |  |  | 1
| 1 | 2 | 01: < 1; 02: < 0.1; 03: < 0.1 |  |  | 12
| 1 | 3 | 01: < 1; 02: < 0.1; 03: < 0.1 |  |  | 16
| 1 | 4 | 01: < 1; 02: < 0.1; 03: < 0.1 |  |  | < 0.1
| 1 | 5 | 01: < 1; 02: < 0.1; 03: < 0.1 |  |  | 18
| 1 | 6 | 01: < 1; 02: < 0.1; 03: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 03
| Analytical results for S in OREAS 24b (Certified Value 0.203 wt.%) | [14 empty] | Certified
S | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab |  |  | 1
| No | No | 01 | 07 | 08 | 09 | 10 | 11 | [7 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | [7 empty] | 3
3
| 1 | 1 | 0.25 | 0.22 | 0.26 | 0.16 | 0.17 | 0.19 | [7 empty] | 1
| 1 | 2 | 0.22 | 0.21 | 0.23 | 0.17 | 0.17 | 0.17 | [7 empty] | 4
| 1 | 3 | 0.23 | 0.22 | 0.28 | 0.2 | 0.17 | 0.17 | [7 empty] | 16
| 1 | 4 | 0.24 | 0.23 | 0.24 | 0.18 | 0.18 | 0.17 | [7 empty] | 0.20333
| 1 | 5 | 0.23 | 0.2 | 0.25 | 0.17 | 0.17 | 0.18 | [7 empty] | 52
| 1 | 6 | 0.26 | 0.21 | 0.22 | 0.18 | 0.17 | 0.16
| Mean |  | 0.23833 | 0.215 | 0.24667 | 0.17667 | 0.17167 | 0.17333
| Median |  | 0.235 | 0.215 | 0.245 | 0.175 | 0.17 | 0.17
| Std Dev. |  | 0.01472 | 0.010488 | 0.021602 | 0.013663 | 0.0040825 | 0.010328
| Rel.Std.Dev. |  | 6.176% | 4.878% | 8.758% | 7.734% | 2.378% | 5.958%
| PDM3 |  | 17.21% | 5.738% | 21.31% | -13.11% | -15.57% | -14.75%
| Z-Score (Absolute) |  | 1.23 | 0.55 | 1.47 | 0.55 | 0.7 | 0.65
| Analytical results for Sb in OREAS 24b (Indicative Value 1.33 ppm) | [14 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 13 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | [7 empty] | 2
2
| 1 | 1 | 2 | 2 | 1 | 0.98 | < 2 | 1 | [7 empty] | 1
| 1 | 2 | 2 | 2 | 0.9 | 0.98 | < 2 | 1.4 | [7 empty] | 7
| 1 | 3 | 2 | 1 | 1.2 | 1.02 | < 2 | 6.8 | [7 empty] | 16
| 1 | 4 | 1 | 1 | 1 | 1.07 | < 2 | 1.5 | [7 empty] | 1.3261
| 1 | 5 | 2 | 2 | 0.9 | 1.05 | 4 | 1.4 | [7 empty] | 13
| 1 | 6 | 2 | 2 | 1.4 | 1.08 | < 2 | 1.5
| Mean |  | 1.8333 | 1.6667 | 1.0667 | 1.03 | 4 | 2.2667
| Median |  | 2 | 2 | 1 | 1.035 | 4 | 1.45
| Std Dev. |  | 0.40825 | 0.5164 | 0.19664 | 0.043818 |  | 2.2286
| Rel.Std.Dev. |  | 22.27% | 30.98% | 18.43% | 4.254% |  | 98.32%
| PDM3 |  | 38.25% | 25.68% | -19.56% | -22.33% | 201.6% | 70.93%
| Z-Score (Absolute) |  | 0.44 | 0.15 | 0.91 | 0.97 | 0.15 | 1.2
| Analytical results for Sc in OREAS 24b (Certified Value 14.1 ppm) | [14 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 06 | 11 | 13 | 14 | [7 empty] | ppm
|  |  | PF*MS | BF*OES | PF*OES/MS | BF*OES | BF*OES | BF*OES/MS | [7 empty] | 1
2
| 1 | 1 | 20 | 14 | 13.3206 | 14 | 15 | 15 | [7 empty] | 1
| 1 | 2 | 20 | 14 | 13.1149 | 14 | 15 | 14 | [7 empty] | 0
| 1 | 3 | 20 | 14 | 13.4692 | 13 | 15 | 15 | [7 empty] | 16
| 1 | 4 | 20 | 14 | 13.4915 | 14 | 15 | 14 | [7 empty] | 14.1229
| 1 | 5 | 20 | 14 | 13.0878 | 14 | 15 | 15 | [7 empty] | 53
| 1 | 6 | 20 | 14 | 13.2035 | 13 | 15 | 15
| Mean |  | 20 | 14 | 13.2812 | 13.6667 | 15 | 14.6667
| Median |  | 20 | 14 | 13.262 | 14 | 15 | 15
| Std Dev. |  | 0 | 0 | 0.1745 | 0.5164 | 0 | 0.5164
| Rel.Std.Dev. |  | 0% | 0% | 1.314% | 3.779% | 0% | 3.521%
| PDM3 |  | 41.61% | -0.8703% | -5.96% | -3.231% | 6.21% | 3.85%
| Z-Score (Absolute) |  | 5.73 | 0.34 | 1.06 | 0.67 | 0.67 | 0.34
| Analytical results for Se in OREAS 24b (Indicative Value < 10 ppm) | [14 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 05 | 07 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | PF*MS | [9 empty] | 1
1
| 1 | 1 | < 50 | < 10 | < 20 | < 0.8 | [9 empty] | 1
| 1 | 2 | < 50 | < 10 | < 20 | < 0.8 | [9 empty] | 8
| 1 | 3 | < 50 | < 10 | < 20 | < 0.8 | [9 empty] | 16
| 1 | 4 | < 50 | < 10 | < 20 | < 0.8 | [9 empty] | < 10
| 1 | 5 | < 50 | < 10 | < 20 | < 0.8 | [9 empty] | 14
| 1 | 6 | < 50 | < 10 | < 20 | < 0.8
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 2.46 | 0.35 | 0.35 | 1
| Analytical results for Si in OREAS 24b (Certified Value 31.12 wt.%) | [14 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 07 | 08 | 10 | 11 | 12 | 13 | 14 |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 29.12 | 31.7 | 30.572 | 29.7 | 31.27 | 32.44 | 33.821 | 29.1039 | 30.853 | 30.867 |  |  |  | 1
| 1 | 2 | 29.22 | 31.7 | 30.217 | 30.1 | 31.23 | 32.44 | 33.12 | 29.0539 | 30.974 | 31.114 |  |  |  | 0
| 1 | 3 | 29.17 | 30.8 | 30.385 | 30.6 | 31.65 | 32.63 | 31.105 | 29.3371 | 31.189 | 31.171 |  |  |  | 16
| 1 | 4 | 29.12 | 31.7 | 30.53 | 31.6 | 31.27 | 32.91 | 31.451 | 29.2904 | 31.026 | 31.072 |  |  |  | 31.1209
| 1 | 5 | 29.03 | 31.3 | 30.698 | 30.8 | 31.79 | 32.26 | 33.265 | 29.2567 | 31.157 | 31.002 |  |  |  | 54
| 1 | 6 | 28.94 | 31.7 | 30.572 | 30.7 | 31.13 | 32.4 | 31.213 | 29.1286 | 31.1 | 31.068
| Mean |  | 29.1 | 31.4833 | 30.4957 | 30.5833 | 31.39 | 32.5133 | 32.3292 | 29.1951 | 31.0498 | 31.049
| Median |  | 29.12 | 31.7 | 30.551 | 30.65 | 31.27 | 32.44 | 32.2855 | 29.1927 | 31.063 | 31.07
| Std Dev. |  | 0.1006 | 0.37103 | 0.16951 | 0.64936 | 0.26442 | 0.22748 | 1.2035 | 0.11465 | 0.12524 | 0.1052
| Rel.Std.Dev. |  | 0.3457% | 1.179% | 0.5559% | 2.123% | 0.8424% | 0.6996% | 3.723% | 0.3927% | 0.4034% | 0.3388%
| PDM3 |  | -6.494% | 1.165% | -2.009% | -1.727% | 0.8648% | 4.474% | 3.883% | -6.188% | -0.2283% | -0.2309%
| Z-Score (Absolute) |  | 2.58 | 0.57 | 0.73 | 0.62 | 0.45 | 1.93 | 1.69 | 2.45 | 0 | 0
| Analytical results for Sm in OREAS 24b (Certified Value 7.17 ppm) | [14 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 7.07 | 7 | 7.5 | 7.5383 | 8.9 | 6.8 | 7.16 | 7.2774 | 7.5 | 7.11 |  |  |  | 1
| 1 | 2 | 7.13 | 6.5 | 7.4 | 7.4752 | 6.8 | 7.04 | 7.02 | 7.2178 | 7.3 | 6.92 |  |  |  | 15
| 1 | 3 | 7.16 | 7 | 7.4 | 7.4483 | 7.4 | 7.5 | 7.43 | 7.3597 | 7.4 | 6.76 |  |  |  | 16
| 1 | 4 | 6.95 | 7 | 7.1 | 7.3177 | 6.8 | 7.27 | 7.41 | 7.1105 | 7.2 | 6.68 |  |  |  | 7.1738
| 1 | 5 | 7.45 | 7.5 | 7.8 | 7.4349 | 7.2 | 6.72 | 7.31 | 7.1149 | 7.2 | 7.13 |  |  |  | 55
| 1 | 6 | 7.18 | 6.5 | 8.2 | 7.6585 | 6.8 | 6.98 | 7.35 | 7.1546 | 7.3 | 6.75
| Mean |  | 7.1567 | 6.9167 | 7.5667 | 7.4788 | 7.3167 | 7.0517 | 7.28 | 7.2058 | 7.3167 | 6.8917
| Median |  | 7.145 | 7 | 7.45 | 7.4618 | 7 | 7.01 | 7.33 | 7.1862 | 7.3 | 6.84
| Std Dev. |  | 0.16585 | 0.37639 | 0.38297 | 0.11372 | 0.81588 | 0.29233 | 0.1595 | 0.098937 | 0.1169 | 0.19364
| Rel.Std.Dev. |  | 2.317% | 5.442% | 5.061% | 1.521% | 11.15% | 4.146% | 2.191% | 1.373% | 1.598% | 2.81%
| PDM3 |  | -0.2388% | -3.584% | 5.476% | 4.252% | 1.992% | -1.702% | 1.48% | 0.4463% | 1.992% | -3.933%
| Z-Score (Absolute) |  | 0.42 | 1.59 | 1.57 | 1.15 | 0.36 | 0.93 | 0.18 | 0.18 | 0.36 | 1.71
| Analytical results for Sn in OREAS 24b (Certified Value 4.65 ppm) | [14 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  | 2
2
| 1 | 1 | 4 | < 10 | < 10 | 5 | 4.9798 | 4.6 | < 50 | 3 | 5 | 5.3749 | 4 | 4 |  | 1
| 1 | 2 | 4 | < 10 | < 10 | 5 | 4.9495 | 4.3 | 480 | 4 | 5 | 5.2059 | 5 | 4 |  | 0
| 1 | 3 | 5 | < 10 | < 10 | 5 | 5.17 | 3.8 | 60 | 4 | 5 | 5.5836 | 5 | 4 |  | 16
| 1 | 4 | 4 | < 10 | < 10 | 5 | 5.4266 | 4.6 | 70 | 4 | 5 | 5.3324 | 5 | 4 |  | 4.6535
| 1 | 5 | 5 | < 10 | < 10 | 5 | 4.945 | 11.3 | 70 | 3 | 5 | 5.5908 | 5 | 4 |  | 56
| 1 | 6 | 5 | < 10 | < 10 | 6 | 4.9765 | 3.4 | 90 | 4 | 5 | 5.9149 | 5 | 4
| Mean |  | 4.5 |  |  | 5.1667 | 5.0746 | 5.3333 | 154 | 3.6667 | 5 | 5.5004 | 4.8333 | 4
| Median |  | 4.5 |  |  | 5 | 4.9782 | 4.45 | 70 | 4 | 5 | 5.4792 | 5 | 4
| Std Dev. |  | 0.54772 |  |  | 0.40825 | 0.19184 | 2.9609 | 182.5651 | 0.5164 | 0 | 0.25205 | 0.40825 | 0
| Rel.Std.Dev. |  | 12.17% |  |  | 7.902% | 3.78% | 55.52% | 118.5% | 14.08% | 0% | 4.582% | 8.447% | 0%
| PDM3 |  | -3.299% |  |  | 11.03% | 9.048% | 14.61% | 3209% | -21.21% | 7.446% | 18.2% | 3.864% | -14.04%
| Z-Score (Absolute) |  | 1.35 | 0 | 0 | 0.45 | 0.2 | 0.9 | 343.9 | 3.6 | 0 | 1.35 | 0.45 | 2.7
| Analytical results for Sr in OREAS 24b (Certified Value 125 ppm) | [14 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*OES | BF*OES | BF*OES/MS |  |  | 0
0
| 1 | 1 | 128 | 124 | 122 | 129 | 111.77 | 134 | 123.5 | 128 | 119.9797 | 119 | 131 |  |  | 1
| 1 | 2 | 129.5 | 126 | 123 | 128 | 111.54 | 119 | 124 | 125.8 | 120.3893 | 123 | 125.2 |  |  | 5
| 1 | 3 | 128.5 | 123 | 124 | 131 | 112.86 | 116 | 124.5 | 129.2 | 121.8619 | 124 | 128.3 |  |  | 16
| 1 | 4 | 126.5 | 122 | 125 | 129 | 111.66 | 123 | 125 | 130.5 | 121.2954 | 121 | 123.2 |  |  | 124.6251
| 1 | 5 | 133.5 | 121 | 122 | 130 | 108.58 | 119 | 125.5 | 129.1 | 121.1538 | 124 | 126.9 |  |  | 57
| 1 | 6 | 130 | 118 | 122 | 133 | 110.5 | 117 | 124 | 128.3 | 119.6281 | 123 | 122.4
| Mean |  | 129.3333 | 122.3333 | 123 | 130 | 111.1517 | 121.3333 | 124.4167 | 128.4833 | 120.7181 | 122.3333 | 126.1667
| Median |  | 129 | 122.5 | 122.5 | 129.5 | 111.6 | 119 | 124.25 | 128.7 | 120.7716 | 123 | 126.05
| Std Dev. |  | 2.3805 | 2.7325 | 1.2649 | 1.7889 | 1.4662 | 6.6533 | 0.73598 | 1.5766 | 0.85707 | 1.9664 | 3.2365
| Rel.Std.Dev. |  | 1.841% | 2.234% | 1.028% | 1.376% | 1.319% | 5.484% | 0.5915% | 1.227% | 0.71% | 1.607% | 2.565%
| PDM3 |  | 3.778% | -1.839% | -1.304% | 4.313% | -10.81% | -2.641% | -0.1673% | 3.096% | -3.135% | -1.839% | 1.237%
| Z-Score (Absolute) |  | 1.87 | 0.2 | 0 | 2.07 | 3.5 | 0.49 | 0.42 | 1.62 | 0.67 | 0.2 | 0.94
| Analytical results for Ta in OREAS 24b (Certified Value 1.32 ppm) | [14 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 2
3
| 1 | 1 | 1.4 | 1 | < 2 | 1.3 | 1.2238 | 1.4 | 1.2 | 1.3 | 1.4442 | 1.3 | 1.4 |  |  | 1
| 1 | 2 | 1.4 | 1 | < 2 | 1.4 | 1.2356 | 1.3 | 1.2 | 1.2 | 1.3888 | 1.3 | 1.2 |  |  | 26
| 1 | 3 | 1.4 | 1 | < 2 | 1.4 | 1.1909 | 1.3 | 1.2 | 1.3 | 1.4389 | 1.3 | 1.2 |  |  | 16
| 1 | 4 | 1.4 | 1 | < 2 | 1.2 | 1.2428 | 1.3 | 1.3 | 1.3 | 1.3989 | 1.3 | 1.4 |  |  | 1.3197
| 1 | 5 | 1.4 | 1 | < 2 | 1.4 | 1.2208 | 1.4 | 1.3 | 1.3 | 1.4168 | 1.3 | 1.4 |  |  | 58
| 1 | 6 | 1.4 | 1 | < 2 | 1.4 | 1.274 | 1.3 | 1.2 | 1.3 | 1.4087 | 1.2 | 1.1
| Mean |  | 1.4 | 1 |  | 1.35 | 1.2313 | 1.3333 | 1.2333 | 1.2833 | 1.4161 | 1.2833 | 1.2833
| Median |  | 1.4 | 1 |  | 1.4 | 1.2297 | 1.3 | 1.2 | 1.3 | 1.4128 | 1.3 | 1.3
| Std Dev. |  | 2.4324e-16 | 0 |  | 0.083666 | 0.02747 | 0.05164 | 0.05164 | 0.040825 | 0.021952 | 0.040825 | 0.13292
| Rel.Std.Dev. |  | 1.737e-14% | 0% |  | 6.197% | 2.231% | 3.873% | 4.187% | 3.181% | 1.55% | 3.181% | 10.36%
| PDM3 |  | 6.084% | -24.23% |  | 2.295% | -6.698% | 1.032% | -6.545% | -2.756% | 7.301% | -2.756% | -2.756%
| Z-Score (Absolute) |  | 1.51 | 3.67 | 3.67 | 0.86 | 0.67 | 0.65 | 0.65 | 0 | 1.72 | 0 | 0
| Analytical results for Tb in OREAS 24b (Certified Value 0.98 ppm) | [14 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.99 | 1.2 | 0.94 | 1.0585 | 1 | 0.99 | 0.9 | 0.98594 | 1 | 1.01 |  |  |  | 1
| 1 | 2 | 0.96 | 1.2 | 0.92 | 1.0396 | 1 | 0.97 | 0.92 | 1.0055 | 0.9 | 0.99 |  |  |  | 16
| 1 | 3 | 0.98 | 1 | 0.93 | 1.0398 | 1 | 1.01 | 0.97 | 1.0078 | 0.9 | 0.99 |  |  |  | 16
| 1 | 4 | 0.97 | 1 | 0.98 | 1.031 | 0.9 | 0.99 | 0.92 | 0.99345 | 0.9 | 0.98 |  |  |  | 0.97969
| 1 | 5 | 0.99 | 1.2 | 1.08 | 1.0203 | 1 | 0.92 | 0.95 | 1.0284 | 0.9 | 1 |  |  |  | 59
| 1 | 6 | 0.99 | 1 | 1.2 | 1.0431 | 0.9 | 0.95 | 0.86 | 0.9917 | 0.9 | 0.96
| Mean |  | 0.98 | 1.1 | 1.0083 | 1.0387 | 0.96667 | 0.97167 | 0.92 | 1.0021 | 0.91667 | 0.98833
| Median |  | 0.985 | 1.1 | 0.96 | 1.0397 | 1 | 0.98 | 0.92 | 0.99948 | 0.9 | 0.99
| Std Dev. |  | 0.012649 | 0.10954 | 0.11071 | 0.012742 | 0.05164 | 0.032506 | 0.038471 | 0.015379 | 0.040825 | 0.017224
| Rel.Std.Dev. |  | 1.291% | 9.959% | 10.98% | 1.227% | 5.342% | 3.345% | 4.182% | 1.535% | 4.454% | 1.743%
| PDM3 |  | 0.03195% | 12.28% | 2.924% | 6.022% | -1.329% | -0.8187% | -6.092% | 2.292% | -6.433% | 0.8826%
| Z-Score (Absolute) |  | 0.13 | 3.71 | 0.77 | 1.74 | 0.56 | 0.4 | 2.05 | 0.58 | 2.16 | 0.13
| Analytical results for Te in OREAS 24b (Indicative Value < 6 ppm) | [14 empty] | Indicative
Te | Round | Replicate | 01: Lab; 02: Lab; 05: Lab; 06: Lab |  |  |  | 1
| No | No | 01: 02; 02: 03; 05: 05; 06: 07 |  |  |  | ppm
|  |  | 01: PF*MS; 02: BF*MS; 05: PF*MS; 06: PF*MS |  |  |  | 2
2
| 1 | 1 | 01: < 50; 02: < 5; 05: < 1; 06: < 6 |  |  |  | 1
| 1 | 2 | 01: < 50; 02: < 5; 05: < 1; 06: < 6 |  |  |  | 9
| 1 | 3 | 01: < 50; 02: < 5; 05: < 1; 06: < 6 |  |  |  | 16
| 1 | 4 | 01: < 50; 02: < 5; 05: < 1; 06: < 6 |  |  |  | < 6
| 1 | 5 | 01: < 50; 02: < 5; 05: < 1; 06: < 6 |  |  |  | 15
| 1 | 6 | 01: < 50; 02: < 5; 05: 2; 06: < 6
| Mean |  | 05: 2
| Median |  | 05: 2
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 01: 13.34; 02: 0.15; 05: 1.2; 06: 0.15
| Analytical results for Th in OREAS 24b (Certified Value 16.5 ppm) | [14 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 1
2
| 1 | 1 | 17.15 | 15.5 | 15.9 | 15.9 | 17.4758 | 16.6 | 17 | 16.88 | 17.5961 | 16.4 | 16.4 |  |  | 1
| 1 | 2 | 17.2 | 16.5 | 15.7 | 15.9 | 16.9792 | 15.8 | 17.35 | 16.72 | 17.4403 | 16.6 | 16.9 |  |  | 0
| 1 | 3 | 17.1 | 16.5 | 15.7 | 15.5 | 16.6581 | 16.3 | 17.05 | 16.98 | 17.4506 | 16.3 | 16.5 |  |  | 16
| 1 | 4 | 16.85 | 15.5 | 16 | 15.6 | 16.5784 | 16.2 | 18.1 | 16.85 | 17.588 | 16.1 | 16.1 |  |  | 16.5355
| 1 | 5 | 17.25 | 16.5 | 16 | 15.9 | 16.5969 | 16.2 | 16.35 | 16.92 | 17.5744 | 16.4 | 16.2 |  |  | 60
| 1 | 6 | 17.15 | 15 | 15.7 | 16.2 | 16.9159 | 16 | 17.1 | 16.71 | 17.3786 | 16.4 | 15.5
| Mean |  | 17.1167 | 15.9167 | 15.8333 | 15.8333 | 16.8674 | 16.1833 | 17.1583 | 16.8433 | 17.5047 | 16.3667 | 16.2667
| Median |  | 17.15 | 16 | 15.8 | 15.9 | 16.787 | 16.2 | 17.075 | 16.865 | 17.5125 | 16.4 | 16.3
| Std Dev. |  | 0.14024 | 0.66458 | 0.15055 | 0.25033 | 0.3422 | 0.27142 | 0.5687 | 0.10857 | 0.092861 | 0.1633 | 0.46762
| Rel.Std.Dev. |  | 0.8193% | 4.175% | 0.9509% | 1.581% | 2.029% | 1.677% | 3.314% | 0.6446% | 0.5305% | 0.9978% | 2.875%
| PDM3 |  | 3.515% | -3.742% | -4.246% | -4.246% | 2.007% | -2.13% | 3.767% | 1.862% | 5.861% | -1.021% | -1.626%
| Z-Score (Absolute) |  | 1.01 | 0.61 | 0.72 | 0.72 | 0.67 | 0.25 | 1.07 | 0.64 | 1.53 | 0 | 0.13
| Analytical results for Ti in OREAS 24b (Certified Value 0.481 wt.%) | [14 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES/MS | 3
3
| 1 | 1 | 0.468 | 0.5 | 0.4706 | 0.41 | 0.509 | 0.49 | 0.504 | 0.484 | 0.486 | 0.492 | 0.4688 | 0.4814 | 0.462 | 1
| 1 | 2 | 0.468 | 0.51 | 0.4712 | 0.46 | 0.496 | 0.49 | 0.504 | 0.471 | 0.486 | 0.486 | 0.4718 | 0.491 | 0.468 | 0
| 1 | 3 | 0.468 | 0.49 | 0.4724 | 0.46 | 0.493 | 0.5 | 0.51 | 0.474 | 0.492 | 0.456 | 0.482 | 0.4916 | 0.468 | 16
| 1 | 4 | 0.468 | 0.5 | 0.4766 | 0.43 | 0.499 | 0.5 | 0.498 | 0.474 | 0.48 | 0.462 | 0.4814 | 0.4832 | 0.468 | 0.48057
| 1 | 5 | 0.468 | 0.49 | 0.473 | 0.43 | 0.495 | 0.49 | 0.51 | 0.484 | 0.48 | 0.48 | 0.476 | 0.5036 | 0.468 | 61
| 1 | 6 | 0.462 | 0.5 | 0.4772 | 0.45 | 0.494 | 0.5 | 0.504 | 0.472 | 0.48 | 0.456 | 0.4784 | 0.4958 | 0.468
| Mean |  | 0.467 | 0.49833 | 0.4735 | 0.44 | 0.49767 | 0.495 | 0.505 | 0.4765 | 0.484 | 0.472 | 0.4764 | 0.4911 | 0.467
| Median |  | 0.468 | 0.5 | 0.4727 | 0.44 | 0.4955 | 0.495 | 0.504 | 0.474 | 0.483 | 0.471 | 0.4772 | 0.4913 | 0.468
| Std Dev. |  | 0.0024495 | 0.0075277 | 0.0027734 | 0.02 | 0.0059217 | 0.0054772 | 0.0045166 | 0.0059245 | 0.004899 | 0.01595 | 0.0052809 | 0.0081866 | 0.0024495
| Rel.Std.Dev. |  | 0.5245% | 1.511% | 0.5857% | 4.545% | 1.19% | 1.107% | 0.8944% | 1.243% | 1.012% | 3.379% | 1.109% | 1.667% | 0.5245%
| PDM3 |  | -2.824% | 3.696% | -1.471% | -8.442% | 3.558% | 3.003% | 5.084% | -0.8468% | 0.7139% | -1.783% | -0.8676% | 2.191% | -2.824%
| Z-Score (Absolute) |  | 0.67 | 1.49 | 0.2 | 2.49 | 1.44 | 1.26 | 1.91 | 0 | 0.48 | 0.33 | 0.01 | 0.99 | 0.67
| Analytical results for Tl in OREAS 24b (Certified Value 0.91 ppm) | [14 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 |  |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS |  |  |  | 2
2
| 1 | 1 | 0.7 | 1 | < 3 | 1 | 0.9244 | 0.9 | < 0.5 | < 0.5 | 0.51453 | 0.9 |  |  |  | 1
| 1 | 2 | 0.7 | < 1 | < 3 | 0.9 | 0.92475 | 0.9 | < 0.5 | < 0.5 | 0.54596 | 1 |  |  |  | 0
| 1 | 3 | 0.8 | 1 | < 3 | 0.9 | 0.89861 | 0.9 | < 0.5 | < 0.5 | 0.56891 | 1 |  |  |  | 16
| 1 | 4 | 0.7 | 1 | < 3 | 0.8 | 0.94605 | 0.9 | < 0.5 | < 0.5 | 0.5515 | 1 |  |  |  | 0.90532
| 1 | 5 | 0.7 | < 1 | < 3 | 0.9 | 0.92671 | 1 | < 0.5 | < 0.5 | 0.5547 | 1 |  |  |  | 62
| 1 | 6 | 0.8 | < 1 | < 3 | 0.9 | 0.87084 | 0.9 | < 0.5 | < 0.5 | 0.61826 | 1
| Mean |  | 0.73333 | 1 |  | 0.9 | 0.91523 | 0.91667 |  |  | 0.55898 | 0.98333
| Median |  | 0.7 | 1 |  | 0.9 | 0.92457 | 0.9 |  |  | 0.5531 | 1
| Std Dev. |  | 0.05164 | 0 |  | 0.063246 | 0.02647 | 0.040825 |  |  | 0.034154 | 0.040825
| Rel.Std.Dev. |  | 7.042% | 0% |  | 7.027% | 2.892% | 4.454% |  |  | 6.11% | 4.152%
| PDM3 |  | -19% | 10.46% |  | -0.5872% | 1.095% | 1.254% |  |  | -38.26% | 8.618%
| Z-Score (Absolute) |  | 0.49 | 0.4 | 3.64 | 0.4 | 0.49 | 0.49 | 3.1 | 3.1 | 1.44 | 0.85
| Analytical results for Tm in OREAS 24b (Certified Value 0.5 ppm) | [14 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.53 | 0.6 | 0.44 | 0.53528 | 0.5 | 0.46 | 0.5 | 0.52671 | 0.49 | 0.51 |  |  |  | 1
| 1 | 2 | 0.53 | 0.6 | 0.42 | 0.52772 | 0.5 | 0.48 | 0.53 | 0.53768 | 0.49 | 0.52 |  |  |  | 18
| 1 | 3 | 0.52 | 0.6 | 0.44 | 0.56958 | 0.5 | 0.47 | 0.53 | 0.52102 | 0.48 | 0.53 |  |  |  | 16
| 1 | 4 | 0.51 | 0.6 | 0.46 | 0.55445 | 0.5 | 0.5 | 0.53 | 0.51685 | 0.49 | 0.48 |  |  |  | 0.50377
| 1 | 5 | 0.52 | 0.6 | 0.51 | 0.53616 | 0.5 | 0.53 | 0.51 | 0.52079 | 0.48 | 0.49 |  |  |  | 63
| 1 | 6 | 0.53 | 0.4 | 0.64 | 0.54012 | 0.5 | 0.48 | 0.47 | 0.50311 | 0.49 | 0.51
| Mean |  | 0.52333 | 0.56667 | 0.485 | 0.54389 | 0.5 | 0.48667 | 0.51167 | 0.52103 | 0.48667 | 0.50667
| Median |  | 0.525 | 0.6 | 0.45 | 0.53814 | 0.5 | 0.48 | 0.52 | 0.5209 | 0.49 | 0.51
| Std Dev. |  | 0.008165 | 0.08165 | 0.081915 | 0.015368 | 0 | 0.025033 | 0.024014 | 0.011386 | 0.005164 | 0.018619
| Rel.Std.Dev. |  | 1.56% | 14.41% | 16.89% | 2.826% | 0% | 5.144% | 4.693% | 2.185% | 1.061% | 3.675%
| PDM3 |  | 3.884% | 12.49% | -3.726% | 7.963% | -0.748% | -3.395% | 1.568% | 3.426% | -3.395% | 0.5754%
| Z-Score (Absolute) |  | 0.52 | 2.11 | 0.89 | 1.28 | 0.34 | 0.83 | 0.09 | 0.44 | 0.83 | 0.09
| Analytical results for U in OREAS 24b (Certified Value 3.31 ppm) | [14 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 2
3
| 1 | 1 | 3.13 | 4 | 2.9 | 3.4 | 3.3468 | 3.7 | 3.32 | 3.33 | 3.5698 | 3.3 | 3.5 |  |  | 1
| 1 | 2 | 3.13 | 4 | 2.8 | 3.2 | 3.2822 | 3.1 | 3.38 | 3.29 | 3.5009 | 3.3 | 3.4 |  |  | 0
| 1 | 3 | 3.14 | 3.5 | 2.8 | 3.2 | 3.2545 | 4 | 3.31 | 3.33 | 3.5318 | 3.3 | 3.4 |  |  | 16
| 1 | 4 | 3.04 | 4 | 2.8 | 3.3 | 3.3367 | 3.3 | 3.48 | 3.37 | 3.4195 | 3.3 | 3.3 |  |  | 3.3137
| 1 | 5 | 3.16 | 4 | 2.9 | 3.1 | 3.2218 | 3.3 | 3.43 | 3.27 | 3.4803 | 3.4 | 3.3 |  |  | 64
| 1 | 6 | 3.07 | 4 | 2.8 | 3.9 | 3.1967 | 3.3 | 3.39 | 3.21 | 3.5415 | 3.3 | 3.2
| Mean |  | 3.1117 | 3.9167 | 2.8333 | 3.35 | 3.2731 | 3.45 | 3.385 | 3.3 | 3.5073 | 3.3167 | 3.35
| Median |  | 3.13 | 4 | 2.8 | 3.25 | 3.2683 | 3.3 | 3.385 | 3.31 | 3.5163 | 3.3 | 3.35
| Std Dev. |  | 0.046224 | 0.20412 | 0.05164 | 0.2881 | 0.060615 | 0.33317 | 0.06473 | 0.056214 | 0.053215 | 0.040825 | 0.10488
| Rel.Std.Dev. |  | 1.486% | 5.212% | 1.823% | 8.6% | 1.852% | 9.657% | 1.912% | 1.703% | 1.517% | 1.231% | 3.131%
| PDM3 |  | -6.098% | 18.19% | -14.5% | 1.094% | -1.227% | 4.112% | 2.15% | -0.4148% | 5.841% | 0.08811% | 1.094%
| Z-Score (Absolute) |  | 2.09 | 4.97 | 4.53 | 0 | 0.67 | 0.88 | 0.31 | 0.44 | 1.38 | 0.29 | 0
| Analytical results for V in OREAS 24b (Certified Value 112 ppm) | [14 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*OES | BF*OES | PF*OES/MS | PF*MS | BF*MS | BF*OES | BF*MS | BF*OES | BF*OES/MS |  |  |  | 0
0
| 1 | 1 | 128 | 100 | 117 | 106.9254 | 113 | 125 | 113 | 106.0897 | 117 | 106 |  |  |  | 1
| 1 | 2 | 128 | 100 | 117 | 105.0594 | 101 | 128 | 113 | 105.2643 | 119 | 102 |  |  |  | 0
| 1 | 3 | 129 | 100 | 118 | 105.9543 | 92 | 123 | 105 | 107.8882 | 122 | 104 |  |  |  | 16
| 1 | 4 | 127 | 100 | 118 | 106.8232 | 106 | 123 | 106 | 106.4034 | 117 | 105 |  |  |  | 112.0198
| 1 | 5 | 130 | 100 | 118 | 104.4262 | 101 | 139 | 110 | 110.3502 | 116 | 105 |  |  |  | 65
| 1 | 6 | 129 | 100 | 120 | 104.6869 | 102 | 132 | 104 | 107.315 | 116 | 105
| Mean |  | 128.5 | 100 | 118 | 105.6459 | 102.5 | 128.3333 | 108.5 | 107.2185 | 117.8333 | 104.5
| Median |  | 128.5 | 100 | 118 | 105.5068 | 101.5 | 126.5 | 108 | 106.8592 | 117 | 105
| Std Dev. |  | 1.0488 | 0 | 1.0954 | 1.0835 | 6.892 | 6.2503 | 4.0373 | 1.7901 | 2.3166 | 1.3784
| Rel.Std.Dev. |  | 0.8162% | 0% | 0.9283% | 1.026% | 6.724% | 4.87% | 3.721% | 1.67% | 1.966% | 1.319%
| PDM3 |  | 14.71% | -10.73% | 5.339% | -5.69% | -8.498% | 14.56% | -3.142% | -4.286% | 5.19% | -6.713%
| Z-Score (Absolute) |  | 2.11 | 0.8 | 1.03 | 0.23 | 0.55 | 2.09 | 0.07 | 0.07 | 1.02 | 0.34
| Analytical results for W in OREAS 24b (Certified Value 4.13 ppm) | [14 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  | 2
2
| 1 | 1 | 5 | < 5 | 4 | 4 | 4.4486 | 4.3 | 8 | 4 | 4 | 4 | 3 |  |  | 1
| 1 | 2 | 5 | < 5 | 4 | 4 | 5.2723 | 3.8 | 8 | 4 | 3.7013 | 4 | 3.7 |  |  | 30
| 1 | 3 | 4 | < 5 | 4 | 4 | 4.326 | 2.7 | 10 | 5 | 3.4125 | 3 | 4 |  |  | 16
| 1 | 4 | 5 | < 5 | 7 | 4 | 3.6014 | 3.8 | 8 | 4 | 3.6536 | 4 | 3.6 |  |  | 4.1331
| 1 | 5 | 4 | < 5 | 4 | 5 | 4.2083 | 3.7 | 8 | 4 | 4.7166 | 4 | 3.7 |  |  | 66
| 1 | 6 | 6 | 5 | 4 | 4 | 4.5812 | 7.2 | 8 | 4 | 3.9046 | 4 | 3.6
| Mean |  | 4.8333 | 5 | 4.5 | 4.1667 | 4.4063 | 4.25 | 8.3333 | 4.1667 | 3.8981 | 3.8333 | 3.6
| Median |  | 5 | 5 | 4 | 4 | 4.3873 | 3.8 | 8 | 4 | 3.8029 | 4 | 3.65
| Std Dev. |  | 0.75277 |  | 1.2247 | 0.40825 | 0.54329 | 1.5372 | 0.8165 | 0.40825 | 0.4505 | 0.40825 | 0.32863
| Rel.Std.Dev. |  | 15.57% |  | 27.22% | 9.798% | 12.33% | 36.17% | 9.798% | 9.798% | 11.56% | 10.65% | 9.129%
| PDM3 |  | 16.94% | 20.97% | 8.877% | 0.812% | 6.61% | 2.828% | 101.6% | 0.812% | -5.686% | -7.253% | -12.9%
| Z-Score (Absolute) |  | 1.35 | 2.53 | 0.67 | 0 | 0.48 | 0.17 | 8.43 | 0 | 0.54 | 0.67 | 1.15
| Analytical results for Y in OREAS 24b (Certified Value 32.5 ppm) | [14 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*OES | BF*OES/MS |  |  | 1
2
| 1 | 1 | 35.3 | 33 | 31 | 31.1 | 33.77 | 33.6 | 33 | 32.6 | 31.2268 | 31 | 31.1 |  |  | 1
| 1 | 2 | 35.6 | 35 | 32 | 31.4 | 33.74 | 32.1 | 33.2 | 32.8 | 31.5579 | 32 | 29.9 |  |  | 0
| 1 | 3 | 35.5 | 35 | 32 | 32.2 | 34.29 | 32.6 | 33.8 | 33.2 | 31.8172 | 32 | 29.4 |  |  | 16
| 1 | 4 | 35 | 34 | 32 | 31.3 | 33.38 | 32.5 | 33.3 | 33.4 | 31.5712 | 31 | 29.4 |  |  | 32.5042
| 1 | 5 | 36.7 | 32 | 32 | 31.4 | 33 | 32 | 33.4 | 33.7 | 31.7888 | 32 | 29.4 |  |  | 67
| 1 | 6 | 35.8 | 32 | 31 | 32.4 | 33.59 | 31.6 | 34.3 | 33 | 31.2873 | 31 | 30
| Mean |  | 35.65 | 33.5 | 31.6667 | 31.6333 | 33.6283 | 32.4 | 33.5 | 33.1167 | 31.5415 | 31.5 | 29.8667
| Median |  | 35.55 | 33.5 | 32 | 31.4 | 33.665 | 32.3 | 33.35 | 33.1 | 31.5645 | 31.5 | 29.65
| Std Dev. |  | 0.58224 | 1.3784 | 0.5164 | 0.53166 | 0.43088 | 0.68993 | 0.47329 | 0.40208 | 0.24575 | 0.54772 | 0.66232
| Rel.Std.Dev. |  | 1.633% | 4.115% | 1.631% | 1.681% | 1.281% | 2.129% | 1.413% | 1.214% | 0.7791% | 1.739% | 2.218%
| PDM3 |  | 9.678% | 3.064% | -2.577% | -2.679% | 3.458% | -0.3207% | 3.064% | 1.884% | -2.962% | -3.09% | -8.115%
| Z-Score (Absolute) |  | 2.44 | 0.82 | 0.55 | 0.57 | 0.92 | 0 | 0.82 | 0.54 | 0.64 | 0.67 | 1.9
| Analytical results for Yb in OREAS 24b (Certified Value 3.24 ppm) | [14 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*MS | BF*MS | BF*MS | BF*MS | BF*OES/MS |  |  |  | 2
3
| 1 | 1 | 3.41 | 3.5 | 3.1 | 3.32 | 3.3 | 2.71 | 3.34 | 3.3018 | 3.3 | 3.18 |  |  |  | 1
| 1 | 2 | 3.41 | 3 | 3.1 | 3.18 | 3.2 | 3.11 | 3.38 | 3.2996 | 3.2 | 3.18 |  |  |  | 19
| 1 | 3 | 3.35 | 3.5 | 3.2 | 3.16 | 3.3 | 3.3 | 3.35 | 3.2023 | 3.2 | 3.24 |  |  |  | 16
| 1 | 4 | 3.27 | 3.5 | 3 | 3.19 | 3.2 | 3.16 | 3.5 | 3.3558 | 3.2 | 3.02 |  |  |  | 3.24
| 1 | 5 | 3.33 | 3.5 | 3.2 | 3.18 | 3.2 | 3.12 | 3.31 | 3.2429 | 3.1 | 3.08 |  |  |  | 68
| 1 | 6 | 3.41 | 3 | 3 | 3.27 | 3.2 | 3.19 | 3.4 | 3.2635 | 3.2 | 3.02
| Mean |  | 3.3633 | 3.3333 | 3.1 | 3.2167 | 3.2333 | 3.0983 | 3.38 | 3.2777 | 3.2 | 3.12
| Median |  | 3.38 | 3.5 | 3.1 | 3.185 | 3.2 | 3.14 | 3.365 | 3.2816 | 3.2 | 3.13
| Std Dev. |  | 0.057504 | 0.2582 | 0.089443 | 0.063456 | 0.05164 | 0.20213 | 0.066633 | 0.053392 | 0.063246 | 0.092952
| Rel.Std.Dev. |  | 1.71% | 7.746% | 2.885% | 1.973% | 1.597% | 6.524% | 1.971% | 1.629% | 1.976% | 2.979%
| PDM3 |  | 3.806% | 2.88% | -4.322% | -0.7212% | -0.2068% | -4.373% | 4.32% | 1.161% | -1.236% | -3.705%
| Z-Score (Absolute) |  | 0.87 | 0.68 | 0.79 | 0.05 | 0.05 | 0.8 | 0.98 | 0.33 | 0.16 | 0.66
| Analytical results for Zn in OREAS 24b (Indicative Value 103 ppm) | [14 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | BF*MS |  |  |  | 0
0
| 1 | 1 | 100 | 150 | 105.5514 | 120 | 100 | 110 | 100 | 65 | 105.4767 | 100 |  |  |  | 1
| 1 | 2 | 100 | 150 | 105.0168 | 110 | 100 | 120 | 100 | 63 | 106.8648 | 100 |  |  |  | 4
| 1 | 3 | 100 | 100 | 105.7893 | 120 | 100 | 100 | 100 | 61 | 105.5439 | 110 |  |  |  | 16
| 1 | 4 | 100 | 150 | 107.7093 | 120 | 100 | 110 | 100 | 62 | 103.1403 | 100 |  |  |  | 102.69
| 1 | 5 | 100 | 100 | 105.5227 | 130 | 100 | 120 | 100 | 65 | 104.8077 | 100 |  |  |  | 16
| 1 | 6 | 100 | 100 | 108.0137 | 120 | 100 | 110 | 100 | 63 | 104.5455 | 100
| Mean |  | 100 | 125 | 106.2672 | 120 | 100 | 111.6667 | 100 | 63.1667 | 105.0631 | 101.6667
| Median |  | 100 | 125 | 105.6703 | 120 | 100 | 110 | 100 | 63 | 105.1422 | 100
| Std Dev. |  | 0 | 27.3861 | 1.264 | 6.3246 | 0 | 7.5277 | 0 | 1.6021 | 1.2392 | 4.0825
| Rel.Std.Dev. |  | 0% | 21.91% | 1.189% | 5.27% | 0% | 6.741% | 0% | 2.536% | 1.18% | 4.016%
| PDM3 |  | -2.62% | 21.73% | 3.483% | 16.86% | -2.62% | 8.741% | -2.62% | -38.49% | 2.311% | -0.9966%
| Z-Score (Absolute) |  | 0.67 | 4.34 | 0.58 | 3.33 | 0.67 | 1.66 | 0.67 | 8.06 | 0.34 | 0.34
| Analytical results for Zr in OREAS 24b (Certified Value 213 ppm) | [14 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 10 | 11 | 12 | 13 | 14 |  |  |  | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | BF*MS | BF*OES | BF*MS | BF*OES | BF*OES/MS |  |  |  | 0
0
| 1 | 1 | 234 | 200 | 199 | 218 | 211.99 | 224 | 222 | 221.38 | 210 | 207.9 |  |  |  | 1
| 1 | 2 | 232 | 220 | 207 | 210 | 214.87 | 228 | 217 | 220.21 | 214 | 204.8 |  |  |  | 0
| 1 | 3 | 229 | 200 | 198 | 217 | 224.52 | 229 | 201 | 214.94 | 212 | 203.3 |  |  |  | 16
| 1 | 4 | 225 | 210 | 203 | 206 | 209.56 | 226 | 203 | 216.5 | 214 | 202 |  |  |  | 213.455
| 1 | 5 | 235 | 200 | 203 | 229 | 212.38 | 220 | 213 | 226.39 | 213 | 208.9 |  |  |  | 69
| 1 | 6 | 233 | 200 | 201 | 214 | 212.14 | 231 | 201 | 211.25 | 211 | 202.6
| Mean |  | 231.3333 | 205 | 201.8333 | 215.6667 | 214.2433 | 226.3333 | 209.5 | 218.445 | 212.3333 | 204.9167
| Median |  | 232.5 | 200 | 202 | 215.5 | 212.26 | 227 | 208 | 218.355 | 212.5 | 204.05
| Std Dev. |  | 3.7238 | 8.3666 | 3.2506 | 7.9162 | 5.3087 | 3.9328 | 9.0719 | 5.3421 | 1.633 | 2.8729
| Rel.Std.Dev. |  | 1.61% | 4.081% | 1.611% | 3.671% | 2.478% | 1.738% | 4.33% | 2.445% | 0.7691% | 1.402%
| PDM3 |  | 8.376% | -3.961% | -5.445% | 1.036% | 0.3693% | 6.033% | -1.853% | 2.338% | -0.5255% | -4%
| Z-Score (Absolute) |  | 1.81 | 0.83 | 1.15 | 0.24 | 0.1 | 1.31 | 0.38 | 0.52 | 0.1 | 0.84
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 24b (Indicative Value 0.127 ppm) | [10 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 3
3
| 1 | 1 | < 0.5 | < 0.5 | 0.3 | < 0.1 | 0.036538 | < 0.05 | 0.08 | < 0.1 |  | 1
| 1 | 2 | < 0.5 | < 0.5 | 0.3 | < 0.1 | 0.032432 | < 0.05 | 0.09 | < 0.1 |  | 7
| 1 | 3 | < 0.5 | < 0.5 | 0.3 | < 0.1 | 0.013793 | < 0.05 | 0.09 | < 0.1 |  | 16
| 1 | 4 | < 0.5 | < 0.5 | 0.3 | 0.2 | 0.021106 | < 0.05 | 0.09 | < 0.1 |  | 0.12684
| 1 | 5 | < 0.5 | < 0.5 | 0.3 | 0.2 | 0.028293 | < 0.05 | 0.08 | < 0.1 |  | 20
| 1 | 6 | < 0.5 | < 0.5 | 0.3 | < 0.1 | 0.01194 | < 0.05 | 0.07 | < 0.1
| Mean |  | 03: 0.3; 05: 0.2; 06: 0.024017; 08: 0.083333
| Median |  | 03: 0.3; 05: 0.2; 06: 0.024699; 08: 0.085
| Std Dev. |  | 03: 0; 05: 0; 06: 0.010047; 08: 0.008165
| Rel.Std.Dev. |  | 03: 0%; 05: 0%; 06: 41.83%; 08: 9.798%
| PDM3 |  | 03: 136.5%; 05: 57.68%; 06: -81.06%; 08: -34.3%
| Z-Score (Absolute) |  | 1.59 | 1.59 | 2.09 | 0.08 | 0.68 | 0.67 | 0.08 | 0.42
| Analytical results for Al in OREAS 24b (Certified Value 8.02 wt.%) | [10 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 8.52 | 7.7493 | 7.3633 | 7.853 | 9.41 | 8.01 | 6.74 |  |  | 1
| 1 | 2 | 8.39 | 7.7577 | 7.9036 | 7.7958 | > 10 | 7.88 | 6.57 |  |  | 0
| 1 | 3 | 8.5 | 7.4922 | 7.8475 | 7.8109 | > 10 | 8.02 | 6.78 |  |  | 16
| 1 | 4 | 8.52 | 7.3707 | 8.0367 | 7.6921 | 8.96 | 7.96 | 6.52 |  |  | 8.0164
| 1 | 5 | 8.31 | 7.2029 | 7.6203 | 7.8595 | 7.49 | 7.74 | 6.91 |  |  | 71
| 1 | 6 | 8.45 | 7.5099 | 7.6632 | 7.7072 | 8.77 | 8.11 | 6.62
| Mean |  | 8.4483 | 7.5138 | 7.7391 | 7.7864 | 8.6575 | 7.9533 | 6.69
| Median |  | 8.475 | 7.5011 | 7.7553 | 7.8034 | 8.865 | 7.985 | 6.68
| Std Dev. |  | 0.084242 | 0.21567 | 0.2402 | 0.071619 | 0.8233 | 0.12894 | 0.14642
| Rel.Std.Dev. |  | 0.9971% | 2.87% | 3.104% | 0.9198% | 9.51% | 1.621% | 2.189%
| PDM3 |  | 5.388% | -6.27% | -3.459% | -2.869% | 7.997% | -0.7868% | -16.55%
| Z-Score (Absolute) |  | 1.64 | 0.67 | 0.12 | 0 | 2.15 | 0.41 | 2.71
| Analytical results for As in OREAS 24b (Indicative Value 8.35 ppm) | [10 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
2
| 1 | 1 | 13 | 8 | 7.5 | 16 | 10 | 8.7923 | 5.7 | 8.9 | 5 | 1
| 1 | 2 | 9 | 8 | 8.1 | 15 | 8 | 7.8984 | 6.5 | 8.2 | 6 | 3
| 1 | 3 | 13 | 7 | 8 | 16 | 12 | 9.3635 | 5.6 | 8.8 | 5 | 16
| 1 | 4 | 16 | 8 | 8 | 16 | 10 | 9.1779 | 5.8 | 8.9 | 6 | 8.35
| 1 | 5 | 14 | 9 | 8.5 | 16 | 9 | 7.9932 | 9.9 | 7.9 | 6 | 21
| 1 | 6 | 9 | 10 | 8.6 | 17 | 9 | 7.9343 | 7.6 | 10.2 | 5
| Mean |  | 12.3333 | 8.3333 | 8.1167 | 16 | 9.6667 | 8.5266 | 6.85 | 8.8167 | 5.5
| Median |  | 13 | 8 | 8.05 | 16 | 9.5 | 8.3927 | 6.15 | 8.85 | 5.5
| Std Dev. |  | 2.8048 | 1.0328 | 0.39707 | 0.63246 | 1.3663 | 0.66712 | 1.6718 | 0.79352 | 0.54772
| Rel.Std.Dev. |  | 22.74% | 12.39% | 4.892% | 3.953% | 14.13% | 7.824% | 24.41% | 9% | 9.959%
| PDM3 |  | 47.7% | -0.1995% | -2.794% | 91.62% | 15.77% | 2.115% | -17.96% | 5.589% | -34.13%
| Z-Score (Absolute) |  | 2.25 | 0.11 | 0.24 | 4.42 | 0.67 | 0 | 0.99 | 0.17 | 1.79
| Analytical results for Ba in OREAS 24b (Certified Value 716 ppm) | [10 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 760 | 728 | 706.1 | 616 | 716 | 703.375 | 659 | 720 | 704 | 1
| 1 | 2 | 750 | 771 | 706.2 | 618 | 721 | 703.8703 | 713 | 710 | 710 | 0
| 1 | 3 | 760 | 755 | 714.7 | 642 | 744 | 690.9064 | 673 | 730 | 683 | 16
| 1 | 4 | 760 | 758 | 703.8 | 618 | 741 | 687.7387 | 664 | 720 | 681 | 715.7515
| 1 | 5 | 750 | 737 | 701.1 | 624 | 736 | 710.4195 | 620 | 700 | 710 | 72
| 1 | 6 | 760 | 779 | 713 | 608 | 717 | 697.8607 | 656 | 730 | 691
| Mean |  | 756.6667 | 754.6667 | 707.4833 | 621 | 729.1667 | 699.0284 | 664.1667 | 718.3333 | 696.5
| Median |  | 760 | 756.5 | 706.15 | 618 | 728.5 | 700.6178 | 661.5 | 720 | 697.5
| Std Dev. |  | 5.164 | 19.4594 | 5.2988 | 11.5065 | 12.6082 | 8.566 | 30.0228 | 11.6905 | 13.2174
| Rel.Std.Dev. |  | 0.6825% | 2.579% | 0.749% | 1.853% | 1.729% | 1.225% | 4.52% | 1.627% | 1.898%
| PDM3 |  | 5.716% | 5.437% | -1.155% | -13.24% | 1.874% | -2.336% | -7.207% | 0.3607% | -2.69%
| Z-Score (Absolute) |  | 1.53 | 1.47 | 0 | 2.69 | 0.67 | 0.26 | 1.35 | 0.34 | 0.34
| Analytical results for Be in OREAS 24b (Certified Value 2.92 ppm) | [10 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
2
| 1 | 1 | 2.6 | 2.6 | 3.4 | 3.7 | 2.6 | 5.34 | 3.2 | 2.79 | 2 | 1
| 1 | 2 | 2.6 | 2.7 | 3.3 | 3.1 | 2.7 | 4.94 | 3.9 | 2.81 | 2 | 25
| 1 | 3 | 2.6 | 2.7 | 3.4 | 3.9 | 2.9 | 4.46 | 3.2 | 2.78 | 3 | 16
| 1 | 4 | 2.6 | 2.5 | 3.3 | 3.8 | 2.8 | 4.56 | 3.4 | 2.7 | 2 | 2.9242
| 1 | 5 | 2.5 | 2.7 | 3.3 | 3.7 | 2.7 | 4.08 | 3.1 | 2.68 | 3 | 73
| 1 | 6 | 2.6 | 2.6 | 3.4 | 3.6 | 2.8 | 3.93 | 3.3 | 2.82 | 2
| Mean |  | 2.5833 | 2.6333 | 3.35 | 3.6333 | 2.75 | 4.5517 | 3.35 | 2.7633 | 2.3333
| Median |  | 2.6 | 2.65 | 3.35 | 3.7 | 2.75 | 4.51 | 3.25 | 2.785 | 2
| Std Dev. |  | 0.040825 | 0.08165 | 0.054772 | 0.28048 | 0.10488 | 0.5272 | 0.2881 | 0.058878 | 0.5164
| Rel.Std.Dev. |  | 1.58% | 3.101% | 1.635% | 7.72% | 3.814% | 11.58% | 8.6% | 2.131% | 22.13%
| PDM3 |  | -11.66% | -9.946% | 14.56% | 24.25% | -5.956% | 55.66% | 14.56% | -5.5% | -20.21%
| Z-Score (Absolute) |  | 0.28 | 0.2 | 0.92 | 1.36 | 0.02 | 2.8 | 0.92 | 0 | 0.67
| Analytical results for Bi in OREAS 24b (Certified Value 0.68 ppm) | [10 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 2 | 0.6 | 0.7 | 0.65 | 0.674 | 0.73 | 0.63 | 0.7 |  | 1
| 1 | 2 | < 2 | 0.6 | 0.6 | 0.72 | 0.58 | 0.78 | 0.67 | 0.7 |  | 10
| 1 | 3 | < 2 | 0.8 | 0.7 | 0.69 | 0.593 | 0.73 | 0.6 | 0.8 |  | 16
| 1 | 4 | < 2 | 0.7 | 0.7 | 0.89 | 0.641 | 0.68 | 0.61 | 1.3 |  | 0.6776
| 1 | 5 | < 2 | 0.6 | 0.6 | 0.7 | 0.638 | 0.9 | 0.7 | 0.7 |  | 74
| 1 | 6 | < 2 | 0.7 | 0.7 | 0.64 | 0.583 | 0.83 | 0.64 | 0.7
| Mean |  | 2 | 0.66667 | 0.66667 | 0.715 | 0.61817 | 0.775 | 0.64167 | 0.81667
| Median |  | 2 | 0.65 | 0.7 | 0.695 | 0.6155 | 0.755 | 0.635 | 0.7
| Std Dev. |  |  | 0.08165 | 0.05164 | 0.09094 | 0.038364 | 0.079687 | 0.037639 | 0.24014
| Rel.Std.Dev. |  |  | 12.25% | 7.746% | 12.72% | 6.206% | 10.28% | 5.866% | 29.4%
| PDM3 |  | 195.2% | -1.613% | -1.613% | 5.52% | -8.771% | 14.38% | -5.302% | 20.52%
| Z-Score (Absolute) |  | 5.27 | 0.27 | 0.27 | 0.27 | 0.8 | 0.93 | 0.54 | 1.39
| Analytical results for Ca in OREAS 24b (Certified Value 1.08 wt.%) | [10 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 1.16 | 1.11 | 1.0982 | 1.0211 | 1.0293 | 1.31 | 1.08 | 0.92 |  | 1
| 1 | 2 | 1.13 | 1.08 | 1.0906 | 1.0911 | 1.0283 | 1.48 | 1.07 | 0.91 |  | 0
| 1 | 3 | 1.15 | 1.07 | 1.0827 | 1.0911 | 1.0261 | 1.3 | 1.09 | 0.94 |  | 16
| 1 | 4 | 1.15 | 1.08 | 1.08 | 1.124 | 1.0147 | 1.33 | 1.08 | 0.89 |  | 1.0838
| 1 | 5 | 1.13 | 1.11 | 1.087 | 1.0542 | 1.0273 | 1.14 | 1.04 | 0.95 |  | 75
| 1 | 6 | 1.15 | 1.1 | 1.0861 | 1.0579 | 1.0135 | 1.32 | 1.09 | 0.91
| Mean |  | 1.145 | 1.0917 | 1.0874 | 1.0732 | 1.0232 | 1.3133 | 1.075 | 0.92
| Median |  | 1.15 | 1.09 | 1.0865 | 1.0745 | 1.0267 | 1.315 | 1.08 | 0.915
| Std Dev. |  | 0.012247 | 0.017224 | 0.0064121 | 0.036141 | 0.0071405 | 0.10801 | 0.018708 | 0.021909
| Rel.Std.Dev. |  | 1.07% | 1.578% | 0.5897% | 3.368% | 0.6979% | 8.224% | 1.74% | 2.381%
| PDM3 |  | 5.651% | 0.7299% | 0.3393% | -0.971% | -5.587% | 21.18% | -0.808% | -15.11%
| Z-Score (Absolute) |  | 1.26 | 0.21 | 0.12 | 0.16 | 1.14 | 4.57 | 0.12 | 3.18
| Analytical results for Cd in OREAS 24b (Indicative Value 0.049 ppm) | [10 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 3
3
| 1 | 1 | < 0.5 | < 0.5 | 0.3 | < 0.02 | < 0.05 | 0.047115 | < 0.1 | 0.05 | < 0.1 | 1
| 1 | 2 | < 0.5 | < 0.5 | 0.3 | < 0.02 | 0.06 | 0.055135 | < 0.1 | 0.04 | 0.2 | 11
| 1 | 3 | < 0.5 | < 0.5 | 0.3 | < 0.02 | < 0.05 | 0.068966 | < 0.1 | 0.05 | < 0.1 | 16
| 1 | 4 | < 0.5 | < 0.5 | 0.3 | < 0.02 | < 0.05 | 0.046231 | < 0.1 | 0.05 | 0.2 | 0.049312
| 1 | 5 | < 0.5 | < 0.5 | 0.3 | < 0.02 | 0.07 | 0.065366 | < 0.1 | 0.05 | < 0.1 | 22
| 1 | 6 | < 0.5 | < 0.5 | 0.3 | < 0.02 | 0.06 | 0.039801 | < 0.1 | 0.05 | < 0.1
| Mean |  | 03: 0.3; 05: 0.063333; 06: 0.053769; 08: 0.048333; 09: 0.2
| Median |  | 03: 0.3; 05: 0.06; 06: 0.051125; 08: 0.05; 09: 0.2
| Std Dev. |  | 03: 0; 05: 0.0057735; 06: 0.011519; 08: 0.0040825; 09: 0
| Rel.Std.Dev. |  | 03: 0%; 05: 9.116%; 06: 21.42%; 08: 8.447%; 09: 0%
| PDM3 |  | 03: 508.4%; 05: 28.43%; 06: 9.039%; 08: -1.984%; 09: 305.6%
| Z-Score (Absolute) |  | 3.02 | 3.02 | 3.79 | 0.67 | 0.15 | 0 | 0.06 | 0.08 | 0.71
| Analytical results for Ce in OREAS 24b (Certified Value 84 ppm) | [10 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 06 | 07 | 08 | 09 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
1
| 1 | 1 | 85.1 | 81.2 | 87.7 | 83.765 | 80.5 | 90.3 | 84 |  |  | 1
| 1 | 2 | 87 | 79.2 | 81.9 | 83.886 | 85.8 | 88 | 84 |  |  | 27
| 1 | 3 | 87.1 | 80.2 | 81 | 84.95 | 82 | 90.2 | 80 |  |  | 16
| 1 | 4 | 84.8 | 80.3 | 81.7 | 84.308 | 79.8 | 89.2 | 80 |  |  | 83.6617
| 1 | 5 | 85.9 | 77.7 | 81.9 | 84.515 | 85.6 | 87.9 | 84 |  |  | 76
| 1 | 6 | 85.3 | 81 | 82.2 | 83.826 | 79 | 93 | 80
| Mean |  | 85.8667 | 79.9333 | 82.7333 | 84.2083 | 82.1167 | 89.7667 | 82
| Median |  | 85.6 | 80.25 | 81.9 | 84.097 | 81.25 | 89.7 | 82
| Std Dev. |  | 0.98522 | 1.3018 | 2.4663 | 0.46918 | 2.9465 | 1.8896 | 2.1909
| Rel.Std.Dev. |  | 1.147% | 1.629% | 2.981% | 0.5572% | 3.588% | 2.105% | 2.672%
| PDM3 |  | 2.636% | -4.456% | -1.11% | 0.6534% | -1.847% | 7.297% | -1.986%
| Z-Score (Absolute) |  | 1.43 | 1.28 | 0 | 0.67 | 0.28 | 3.22 | 0.34
| Analytical results for Co in OREAS 24b (Certified Value 16.9 ppm) | [10 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 14 | 15 | 17.1 | 17.9 | 16.2 | 18.6462 | 16.8 | 16.9 | 15.9 | 1
| 1 | 2 | 15 | 15 | 17.2 | 16.3 | 16.1 | 18.4962 | 18.5 | 16.3 | 15.9 | 5
| 1 | 3 | 15 | 15 | 17.7 | 18.4 | 16.7 | 18.4631 | 16.7 | 16.8 | 16.2 | 16
| 1 | 4 | 14 | 20 | 17.5 | 15.9 | 16.8 | 17.8935 | 17.6 | 16.6 | 15.6 | 16.876
| 1 | 5 | 14 | 20 | 17.7 | 15.8 | 16.4 | 17.3122 | 16.5 | 16 | 15.7 | 77
| 1 | 6 | 15 | 15 | 18 | 16 | 16.1 | 17.0378 | 17.4 | 17.2 | 15.8
| Mean |  | 14.5 | 16.6667 | 17.5333 | 16.7167 | 16.3833 | 17.9748 | 17.25 | 16.6333 | 15.85
| Median |  | 14.5 | 15 | 17.6 | 16.15 | 16.3 | 18.1783 | 17.1 | 16.7 | 15.85
| Std Dev. |  | 0.54772 | 2.582 | 0.33862 | 1.1339 | 0.30605 | 0.67596 | 0.74498 | 0.43205 | 0.20736
| Rel.Std.Dev. |  | 3.777% | 15.49% | 1.931% | 6.783% | 1.868% | 3.761% | 4.319% | 2.597% | 1.308%
| PDM3 |  | -14.08% | -1.241% | 3.895% | -0.9443% | -2.919% | 6.511% | 2.216% | -1.438% | -6.08%
| Z-Score (Absolute) |  | 2.5 | 0 | 1 | 0.06 | 0.33 | 1.51 | 0.67 | 0.04 | 0.94
| Analytical results for Cr in OREAS 24b (Certified Value 118 ppm) | [10 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 129 | 107 | 116 | 109 | 83.7639 | 112 | 128 | 108 |  | 1
| 1 | 2 | 127 | 104 | 115 | 119 | 90.6499 | 133 | 123 | 105 |  | 0
| 1 | 3 | 131 | 104 | 108 | 121 | 79.6562 | 120 | 128 | 111 |  | 16
| 1 | 4 | 132 | 110 | 112 | 127 | 109.9042 | 122 | 126 | 111 |  | 118.2905
| 1 | 5 | 134 | 102 | 119 | 119 | 85.0982 | 139 | 123 | 114 |  | 78
| 1 | 6 | 129 | 104 | 109 | 119 | 82.6596 | 128 | 126 | 111
| Mean |  | 130.3333 | 105.1667 | 113.1667 | 119 | 88.622 | 125.6667 | 125.6667 | 110
| Median |  | 130 | 104 | 113.5 | 119 | 84.4311 | 125 | 126 | 111
| Std Dev. |  | 2.5033 | 2.8577 | 4.2622 | 5.7966 | 11.0363 | 9.6885 | 2.2509 | 3.0984
| Rel.Std.Dev. |  | 1.921% | 2.717% | 3.766% | 4.871% | 12.45% | 7.71% | 1.791% | 2.817%
| PDM3 |  | 10.18% | -11.09% | -4.332% | 0.5998% | -25.08% | 6.236% | 6.236% | -7.009%
| Z-Score (Absolute) |  | 1 | 0.77 | 0.21 | 0.21 | 1.93 | 0.67 | 0.67 | 0.43
| Analytical results for Cs in OREAS 24b (Certified Value 10.7 ppm) | [10 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 10.3 | 10.6 | 10.1 | 10.6 | 12.199 | 10.7 | 11.15 |  |  | 1
| 1 | 2 | 11 | 10.5 | 9.26 | 10.7 | 12.1297 | 11.4 | 11.1 |  |  | 0
| 1 | 3 | 10.7 | 10.5 | 9.89 | 11 | 12.2473 | 11 | 11.25 |  |  | 16
| 1 | 4 | 10.2 | 10.3 | 10.4 | 10.9 | 12.1367 | 10.8 | 10.95 |  |  | 10.6533
| 1 | 5 | 10.3 | 10.2 | 10.1 | 10.8 | 11.9893 | 9.77 | 11.05 |  |  | 79
| 1 | 6 | 10.5 | 10.5 | 10.1 | 10.4 | 11.9592 | 10.9 | 11.55
| Mean |  | 10.5 | 10.4333 | 9.975 | 10.7333 | 12.1102 | 10.7617 | 11.175
| Median |  | 10.4 | 10.5 | 10.1 | 10.75 | 12.1332 | 10.85 | 11.125
| Std Dev. |  | 0.30332 | 0.15055 | 0.38625 | 0.21602 | 0.11419 | 0.5426 | 0.20917
| Rel.Std.Dev. |  | 2.889% | 1.443% | 3.872% | 2.013% | 0.9429% | 5.042% | 1.872%
| PDM3 |  | -1.439% | -2.065% | -6.367% | 0.7515% | 13.68% | 1.017% | 4.897%
| Z-Score (Absolute) |  | 0.52 | 0.67 | 1.7 | 0 | 3.09 | 0.06 | 0.99
| Analytical results for Cu in OREAS 24b (Certified Value 38 ppm) | [10 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 40 | 40 | 40 | 27.5 | 37 | 37.2413 | 44.4 | 38.7 | 36.3 | 1
| 1 | 2 | 38 | 36 | 38 | 27.9 | 40 | 36.2789 | 42.4 | 37.3 | 36.2 | 12
| 1 | 3 | 39 | 36 | 38 | 27.8 | 39 | 35.9754 | 38.9 | 37.8 | 37.9 | 16
| 1 | 4 | 39 | 36 | 39 | 28.3 | 41 | 37.2332 | 39.3 | 36.8 | 35.2 | 37.9595
| 1 | 5 | 38 | 38 | 40 | 29 | 38 | 37.1522 | 39 | 36.7 | 38.3 | 80
| 1 | 6 | 39 | 36 | 40 | 28.5 | 38 | 35.4965 | 38.3 | 38.9 | 35.7
| Mean |  | 38.8333 | 37 | 39.1667 | 28.1667 | 38.8333 | 36.5629 | 40.3833 | 37.7 | 36.6
| Median |  | 39 | 36 | 39.5 | 28.1 | 38.5 | 36.7156 | 39.15 | 37.55 | 36.25
| Std Dev. |  | 0.75277 | 1.6733 | 0.98319 | 0.54283 | 1.472 | 0.75098 | 2.4425 | 0.94021 | 1.2329
| Rel.Std.Dev. |  | 1.938% | 4.522% | 2.51% | 1.927% | 3.79% | 2.054% | 6.048% | 2.494% | 3.369%
| PDM3 |  | 2.302% | -2.528% | 3.18% | -25.8% | 2.302% | -3.679% | 6.385% | -0.6837% | -3.582%
| Z-Score (Absolute) |  | 0.67 | 0.42 | 0.87 | 5.67 | 0.67 | 0.68 | 1.6 | 0 | 0.65
| Analytical results for Dy in OREAS 24b (Indicative Value 4.47 ppm) | [10 empty] | Indicative
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 02; 02: 04; 03: 07 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS | 2
2
| 1 | 1 | 01: 4.6; 02: 4.03; 03: 4.9 | 1
| 1 | 2 | 01: 4.8; 02: 3.77; 03: 4.7 | 17
| 1 | 3 | 01: 4.9; 02: 4.05; 03: 4.8 | 16
| 1 | 4 | 01: 4.55; 02: 4.28; 03: 4.6 | 4.4656
| 1 | 5 | 01: 4.55; 02: 4.06; 03: 4.6 | 23
| 1 | 6 | 01: 4.6; 02: 3.99; 03: 4.6
| Mean |  | 01: 4.6667; 02: 4.03; 03: 4.7
| Median |  | 01: 4.6; 02: 4.04; 03: 4.65
| Std Dev. |  | 01: 0.1472; 02: 0.1631; 03: 0.12649
| Rel.Std.Dev. |  | 01: 3.154%; 02: 4.047%; 03: 2.691%
| PDM3 |  | 01: 4.504%; 02: -9.754%; 03: 5.25%
| Z-Score (Absolute) |  | 01: 0; 02: 12.88; 03: 0.67
| Analytical results for Er in OREAS 24b (Indicative Value 2.54 ppm) | [10 empty] | Indicative
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 02; 02: 04; 03: 07 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS | 2
2
| 1 | 1 | 01: 2.35; 02: 1.87; 03: 2.5 | 1
| 1 | 2 | 01: 3.8; 02: 1.81; 03: 2.6 | 18
| 1 | 3 | 01: 3.05; 02: 1.74; 03: 2.5 | 16
| 1 | 4 | 01: 2.55; 02: 1.96; 03: 2.4 | 2.5383
| 1 | 5 | 01: 2.35; 02: 1.77; 03: 2.6 | 24
| 1 | 6 | 01: 2.5; 02: 1.78; 03: 2.5
| Mean |  | 01: 2.7667; 02: 1.8217; 03: 2.5167
| Median |  | 01: 2.525; 02: 1.795; 03: 2.5
| Std Dev. |  | 01: 0.56804; 02: 0.08085; 03: 0.075277
| Rel.Std.Dev. |  | 01: 20.53%; 02: 4.438%; 03: 2.991%
| PDM3 |  | 01: 8.995%; 02: -28.23%; 03: -0.8536%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.87; 03: 0
| Analytical results for Eu in OREAS 24b (Indicative Value 1.36 ppm) | [10 empty] | Indicative
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 02; 02: 04; 03: 07 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS | 2
2
| 1 | 1 | 01: 1.45; 02: 1.44; 03: 1.33 | 1
| 1 | 2 | 01: 1.5; 02: 1.31; 03: 1.35 | 19
| 1 | 3 | 01: 1.55; 02: 1.26; 03: 1.33 | 16
| 1 | 4 | 01: 1.55; 02: 1.28; 03: 1.24 | 1.3612
| 1 | 5 | 01: 1.45; 02: 1.3; 03: 1.29 | 25
| 1 | 6 | 01: 1.5; 02: 1.26; 03: 1.27
| Mean |  | 01: 1.5; 02: 1.3083; 03: 1.3017
| Median |  | 01: 1.5; 02: 1.29; 03: 1.31
| Std Dev. |  | 01: 0.044721; 02: 0.067651; 03: 0.042151
| Rel.Std.Dev. |  | 01: 2.981%; 02: 5.171%; 03: 3.238%
| PDM3 |  | 01: 10.2%; 02: -3.885%; 03: -4.375%
| Z-Score (Absolute) |  | 01: 19.39; 02: 0; 03: 0.67
| Analytical results for Fe in OREAS 24b (Certified Value 4.39 wt.%) | [10 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 4.9 | 4.4271 | 4.33 | 4.2549 | 4.5 | 4.6 | 4.32 |  |  | 1
| 1 | 2 | 4.82 | 4.3922 | 4.46 | 4.2397 | 4.87 | 4.51 | 4.29 |  |  | 0
| 1 | 3 | 4.92 | 4.3443 | 4.55 | 4.1955 | 4.51 | 4.62 | 4.34 |  |  | 16
| 1 | 4 | 4.91 | 4.351 | 4.53 | 4.1954 | 4.56 | 4.56 | 4.24 |  |  | 4.3941
| 1 | 5 | 4.8 | 4.3172 | 4.39 | 4.2507 | 4.2 | 4.43 | 4.36 |  |  | 81
| 1 | 6 | 4.91 | 4.3657 | 4.47 | 4.1982 | 4.45 | 4.64 | 4.35
| Mean |  | 4.8767 | 4.3662 | 4.455 | 4.2224 | 4.515 | 4.56 | 4.3167
| Median |  | 4.905 | 4.3583 | 4.465 | 4.2189 | 4.505 | 4.58 | 4.33
| Std Dev. |  | 0.052409 | 0.038721 | 0.083367 | 0.028977 | 0.21529 | 0.07874 | 0.045019
| Rel.Std.Dev. |  | 1.075% | 0.8868% | 1.871% | 0.6863% | 4.768% | 1.727% | 1.043%
| PDM3 |  | 10.98% | -0.6327% | 1.387% | -3.907% | 2.753% | 3.777% | -1.761%
| Z-Score (Absolute) |  | 2.71 | 0.57 | 0 | 1.49 | 0.39 | 0.67 | 0.89
| Analytical results for Ga in OREAS 24b (Certified Value 20.1 ppm) | [10 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 20 | 18.8 | 19.3 | 22.9 | 20.5 | 20.52 | 19.1 | 20.7 |  | 1
| 1 | 2 | 20 | 18.8 | 19.3 | 20.5 | 20.5 | 20.14 | 21.6 | 19.8 |  | 0
| 1 | 3 | 20 | 19 | 19.6 | 21 | 21.1 | 20.33 | 20.9 | 20.4 |  | 16
| 1 | 4 | 20 | 18.6 | 19.2 | 21.7 | 21.3 | 20.05 | 20 | 19.75 |  | 20.1212
| 1 | 5 | 20 | 19.2 | 19.4 | 20.4 | 21.3 | 19.57 | 21.1 | 19.35 |  | 82
| 1 | 6 | 20 | 19.4 | 19.7 | 22 | 20.8 | 19.39 | 19.9 | 20.7
| Mean |  | 20 | 18.9667 | 19.4167 | 21.4167 | 20.9167 | 20 | 20.4333 | 20.1167
| Median |  | 20 | 18.9 | 19.35 | 21.35 | 20.95 | 20.095 | 20.45 | 20.1
| Std Dev. |  | 0 | 0.29439 | 0.19408 | 0.96626 | 0.37103 | 0.43781 | 0.92448 | 0.56273
| Rel.Std.Dev. |  | 0% | 1.552% | 0.9995% | 4.512% | 1.774% | 2.189% | 4.524% | 2.797%
| PDM3 |  | -0.6026% | -5.738% | -3.502% | 6.438% | 3.953% | -0.6026% | 1.551% | -0.02278%
| Z-Score (Absolute) |  | 0.08 | 1.45 | 0.85 | 1.8 | 1.14 | 0.08 | 0.5 | 0.08
| Analytical results for Gd in OREAS 24b (Indicative Value 6.02 ppm) | [10 empty] | Indicative
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 07 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS |  | 2
2
| 1 | 1 | 01: 5.8; 02: 6.04; 03: 6.2 |  | 1
| 1 | 2 | 01: 6.2; 02: 5.69; 03: 6 |  | 20
| 1 | 3 | 01: 6.2; 02: 5.93; 03: 6.2 |  | 16
| 1 | 4 | 01: 6.6; 02: 5.98; 03: 6 |  | 6.0244
| 1 | 5 | 01: 6; 02: 5.95; 03: 5.7 |  | 26
| 1 | 6 | 01: 6; 02: 6.05; 03: 5.9
| Mean |  | 01: 6.1333; 02: 5.94; 03: 6
| Median |  | 01: 6.1; 02: 5.965; 03: 6
| Std Dev. |  | 01: 0.27325; 02: 0.13145; 03: 0.18974
| Rel.Std.Dev. |  | 01: 4.455%; 02: 2.213%; 03: 3.162%
| PDM3 |  | 01: 1.807%; 02: -1.402%; 03: -0.4058%
| Z-Score (Absolute) |  | 01: 1.5; 02: 0.67; 03: 0
| Analytical results for Ge in OREAS 24b (Indicative Value 0.83 ppm) | [10 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  | 1
| No | No | 01: 05; 02: 07; 03: 08 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS |  | 2
2
| 1 | 1 | 01: 1.9; 02: 0.5; 03: 0.2 |  | 1
| 1 | 2 | 01: 1.6; 02: 0.6; 03: 0.2 |  | 32
| 1 | 3 | 01: 1.9; 02: 0.5; 03: 0.19 |  | 16
| 1 | 4 | 01: 1.7; 02: 0.5; 03: 0.2 |  | 0.83444
| 1 | 5 | 01: 1.9; 02: 0.5; 03: 0.2 |  | 20
| 1 | 6 | 01: 1.7; 02: 0.8; 03: 0.21
| Mean |  | 01: 1.7833; 02: 0.56667; 03: 0.2
| Median |  | 01: 1.8; 02: 0.5; 03: 0.2
| Std Dev. |  | 01: 0.13292; 02: 0.12111; 03: 0.0063246
| Rel.Std.Dev. |  | 01: 7.453%; 02: 21.37%; 03: 3.162%
| PDM3 |  | 01: 113.7%; 02: -32.09%; 03: -76.03%
| Z-Score (Absolute) |  | 01: 2.24; 02: 0; 03: 0.67
| Analytical results for Hf in OREAS 24b (Certified Value 3.9 ppm) | [10 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | 4.6 | 4.3 | 3.67 | 4.3 | 4.0519 | 2.8 | 3.7 | 3.8 |  | 1
| 1 | 2 | 4.8 | 4.3 | 3.16 | 4.3 | 3.9903 | 3 | 3.6 | 3.8 |  | 33
| 1 | 3 | 4.8 | 4.3 | 3.56 | 4.3 | 4.0502 | 3.2 | 3.7 | 3.6 |  | 16
| 1 | 4 | 4.6 | 4.3 | 3.53 | 4.6 | 4.0975 | 2.7 | 3.7 | 3.5 |  | 3.9008
| 1 | 5 | 4.6 | 4.3 | 3.51 | 4.6 | 4.0176 | 2.8 | 3.6 | 3.6 |  | 83
| 1 | 6 | 4.6 | 4.3 | 3.48 | 4.3 | 4.0299 | 3.2 | 3.7 | 3.5
| Mean |  | 4.6667 | 4.3 | 3.485 | 4.4 | 4.0396 | 2.95 | 3.6667 | 3.6333
| Median |  | 4.6 | 4.3 | 3.52 | 4.3 | 4.04 | 2.9 | 3.7 | 3.6
| Std Dev. |  | 0.10328 | 0 | 0.17213 | 0.15492 | 0.036415 | 0.21679 | 0.05164 | 0.13663
| Rel.Std.Dev. |  | 2.213% | 0% | 4.939% | 3.521% | 0.9015% | 7.349% | 1.408% | 3.76%
| PDM3 |  | 19.63% | 10.23% | -10.66% | 12.8% | 3.558% | -24.37% | -6.002% | -6.856%
| Z-Score (Absolute) |  | 1.35 | 0.74 | 0.61 | 0.9 | 0.31 | 1.49 | 0.31 | 0.36
| Analytical results for Ho in OREAS 24b (Indicative Value 0.8 ppm) | [10 empty] | Indicative
Ho | Round | Replicate | 02: Lab; 03: Lab; 04: Lab |  | 1
| No | No | 02: 02; 03: 04; 04: 07 |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*MS |  | 2
2
| 1 | 1 | 02: 0.88; 03: 0.55; 04: 0.9 |  | 1
| 1 | 2 | 02: 0.96; 03: 0.6; 04: 1 |  | 15
| 1 | 3 | 02: 0.94; 03: 0.66; 04: 0.9 |  | 16
| 1 | 4 | 02: 0.88; 03: 0.63; 04: 0.9 |  | 0.80111
| 1 | 5 | 02: 0.88; 03: 0.59; 04: 0.9 |  | 21
| 1 | 6 | 02: 0.86; 03: 0.59; 04: 0.9
| Mean |  | 02: 0.9; 03: 0.60333; 04: 0.91667
| Median |  | 02: 0.88; 03: 0.595; 04: 0.9
| Std Dev. |  | 02: 0.04; 03: 0.037771; 04: 0.040825
| Rel.Std.Dev. |  | 02: 4.444%; 03: 6.26%; 04: 4.454%
| PDM3 |  | 02: 12.34%; 03: -24.69%; 04: 14.42%
| Z-Score (Absolute) |  | 02: 0; 03: 12; 04: 0.67
| Analytical results for In in OREAS 24b (Indicative Value 0.077 ppm) | [10 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.06 | < 0.5 | < 0.02 | 0.08 | 0.084615 | < 0.1 | 0.08 |  |  | 1
| 1 | 2 | 0.06 | < 0.5 | < 0.02 | 0.08 | 0.08 | < 0.1 | 0.074 |  |  | 8
| 1 | 3 | 0.08 | < 0.5 | < 0.02 | 0.08 | 0.085714 | < 0.1 | 0.078 |  |  | 16
| 1 | 4 | 0.06 | < 0.5 | 0.04 | 0.09 | 0.081407 | < 0.1 | 0.072 |  |  | 0.076718
| 1 | 5 | 0.08 | < 0.5 | < 0.02 | 0.07 | 0.084878 | < 0.1 | 0.077 |  |  | 22
| 1 | 6 | 0.08 | < 0.5 | 0.03 | 0.07 | 0.074627 | < 0.1 | 0.079
| Mean |  | 0.07 |  | 0.035 | 0.078333 | 0.081874 |  | 0.076667
| Median |  | 0.07 |  | 0.035 | 0.08 | 0.083011 |  | 0.0775
| Std Dev. |  | 0.010954 |  | 0.0070711 | 0.0075277 | 0.0041842 |  | 0.0030768
| Rel.Std.Dev. |  | 15.65% |  | 20.2% | 9.61% | 5.111% |  | 4.013%
| PDM3 |  | -8.757% |  | -54.38% | 2.105% | 6.72% |  | -0.06743%
| Z-Score (Absolute) |  | 0.67 | 17.53 | 5.9 | 0.17 | 0.53 | 2.7 | 0
| Analytical results for K in OREAS 24b (Certified Value 2.81 wt.%) | [10 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 2.95 | 2.83 | 2.719 | 2.6 | 2.7505 | 2.868 | 3.2 | 2.86 | 2.75 | 1
| 1 | 2 | 2.9 | 2.76 | 2.7178 | 2.62 | 2.9767 | 2.8245 | 3.51 | 2.86 | 2.77 | 0
| 1 | 3 | 2.93 | 2.73 | 2.6495 | 2.72 | 2.9415 | 2.8307 | 3.31 | 2.93 | 2.91 | 16
| 1 | 4 | 2.94 | 2.75 | 2.6518 | 2.68 | 3.0376 | 2.7995 | 3.29 | 2.84 | 2.74 | 2.8062
| 1 | 5 | 2.88 | 2.79 | 2.6722 | 2.7 | 2.8862 | 2.8657 | 2.71 | 2.8 | 2.87 | 84
| 1 | 6 | 2.95 | 2.8 | 2.664 | 2.61 | 2.902 | 2.8123 | 3.15 | 2.97 | 2.69
| Mean |  | 2.925 | 2.7767 | 2.6791 | 2.655 | 2.9158 | 2.8335 | 3.195 | 2.8767 | 2.7883
| Median |  | 2.935 | 2.775 | 2.6681 | 2.65 | 2.9217 | 2.8276 | 3.245 | 2.86 | 2.76
| Std Dev. |  | 0.02881 | 0.036697 | 0.031581 | 0.051284 | 0.097566 | 0.027982 | 0.26786 | 0.062183 | 0.084004
| Rel.Std.Dev. |  | 0.9849% | 1.322% | 1.179% | 1.932% | 3.346% | 0.9876% | 8.384% | 2.162% | 3.013%
| PDM3 |  | 4.232% | -1.054% | -4.532% | -5.389% | 3.902% | 0.9699% | 13.85% | 2.51% | -0.6381%
| Z-Score (Absolute) |  | 0.75 | 0.47 | 1.27 | 1.46 | 0.67 | 0 | 2.96 | 0.35 | 0.37
| Analytical results for La in OREAS 24b (Certified Value 42.4 ppm) | [10 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 40 | 43.3 | 41.3 | 45.3 | 44.47 | 41.8 | 47 | 40.2 |  | 1
| 1 | 2 | 40 | 42.9 | 41.1 | 42.9 | 44.3 | 43.8 | 45.6 | 39.7 |  | 25
| 1 | 3 | 40 | 42.7 | 41.8 | 40.7 | 45.19 | 42.4 | 46.6 | 39.1 |  | 16
| 1 | 4 | 40 | 42.8 | 40.8 | 42.1 | 44.66 | 42 | 46.1 | 37.9 |  | 42.3548
| 1 | 5 | 40 | 42 | 40.4 | 42.1 | 44.28 | 44.1 | 44.9 | 40.3 |  | 85
| 1 | 6 | 40 | 44.5 | 42.3 | 41.3 | 44.31 | 40.9 | 47.9 | 38.7
| Mean |  | 40 | 43.0333 | 41.2833 | 42.4 | 44.535 | 42.5 | 46.35 | 39.3167
| Median |  | 40 | 42.85 | 41.2 | 42.1 | 44.39 | 42.2 | 46.35 | 39.4
| Std Dev. |  | 0 | 0.83347 | 0.68532 | 1.6087 | 0.35207 | 1.2296 | 1.0597 | 0.93041
| Rel.Std.Dev. |  | 0% | 1.937% | 1.66% | 3.794% | 0.7905% | 2.893% | 2.286% | 2.366%
| PDM3 |  | -5.56% | 1.602% | -2.53% | 0.1067% | 5.147% | 0.3428% | 9.433% | -7.173%
| Z-Score (Absolute) |  | 1.02 | 0.24 | 0.48 | 0.02 | 0.86 | 0.02 | 1.62 | 1.3
| Analytical results for Li in OREAS 24b (Certified Value 52 ppm) | [10 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
1
| 1 | 1 | 46 | 52 | 51 | 49.2 | 52.4269 | 58.7 | 51.6 | 53.6 |  | 1
| 1 | 2 | 45.5 | 52 | 51 | 49.7 | 52.8627 | 69.7 | 51.1 | 51.7 |  | 0
| 1 | 3 | 45.5 | 52 | 54 | 51.2 | 52.1202 | 60.7 | 51.5 | 52.5 |  | 16
| 1 | 4 | 45 | 53 | 53 | 51.3 | 51.8462 | 58.6 | 50.3 | 52.3 |  | 51.7003
| 1 | 5 | 49.5 | 52 | 52 | 51 | 52.5629 | 59 | 50.6 | 52.2 |  | 86
| 1 | 6 | 44 | 53 | 51 | 50.4 | 51.591 | 56.8 | 53 | 48.6
| Mean |  | 45.9167 | 52.3333 | 52 | 50.4667 | 52.235 | 60.5833 | 51.35 | 51.8167
| Median |  | 45.5 | 52 | 51.5 | 50.7 | 52.2736 | 58.85 | 51.3 | 52.25
| Std Dev. |  | 1.8819 | 0.5164 | 1.2649 | 0.86178 | 0.47273 | 4.6353 | 0.95237 | 1.6964
| Rel.Std.Dev. |  | 4.099% | 0.9867% | 2.433% | 1.708% | 0.905% | 7.651% | 1.855% | 3.274%
| PDM3 |  | -11.19% | 1.224% | 0.5797% | -2.386% | 1.034% | 17.18% | -0.6775% | 0.2251%
| Z-Score (Absolute) |  | 8.22 | 0.58 | 0.13 | 1.98 | 0.45 | 11.9 | 0.77 | 0.13
| Analytical results for Lu in OREAS 24b (Indicative Value 0.32 ppm) | [10 empty] | Indicative
Lu | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 02: 02; 03: 04; 04: 06; 05: 07 |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS |  | 2
2
| 1 | 1 | 02: 0.42; 03: 0.27; 04: 0.322; 05: 0.3 |  | 1
| 1 | 2 | 02: 0.4; 03: 0.24; 04: 0.321; 05: 0.3 |  | 17
| 1 | 3 | 02: 0.4; 03: 0.29; 04: 0.327; 05: 0.4 |  | 16
| 1 | 4 | 02: 0.38; 03: 0.29; 04: 0.329; 05: 0.3 |  | 0.32258
| 1 | 5 | 02: 0.36; 03: 0.27; 04: 0.328; 05: 0.3 |  | 23
| 1 | 6 | 02: 0.38; 03: 0.28; 04: 0.335; 05: 0.3
| Mean |  | 02: 0.39; 03: 0.27333; 04: 0.327; 05: 0.31667
| Median |  | 02: 0.39; 03: 0.275; 04: 0.3275; 05: 0.3
| Std Dev. |  | 02: 0.020976; 03: 0.018619; 04: 0.005099; 05: 0.040825
| Rel.Std.Dev. |  | 02: 5.379%; 03: 6.812%; 04: 1.559%; 05: 12.89%
| PDM3 |  | 02: 20.9%; 03: -15.27%; 04: 1.369%; 05: -1.834%
| Z-Score (Absolute) |  | 02: 1.71; 03: 1.22; 04: 0.13; 05: 0.13
| Analytical results for Mg in OREAS 24b (Certified Value 1.65 wt.%) | [10 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.74 | 1.7063 | 1.5988 | 1.6103 | 1.98 | 1.61 | 1.61 |  |  | 1
| 1 | 2 | 1.71 | 1.7021 | 1.6488 | 1.6138 | 2.13 | 1.58 | 1.61 |  |  | 0
| 1 | 3 | 1.74 | 1.6794 | 1.6934 | 1.5991 | 2.14 | 1.62 | 1.66 |  |  | 16
| 1 | 4 | 1.74 | 1.6939 | 1.6819 | 1.5809 | 1.91 | 1.6 | 1.57 |  |  | 1.6479
| 1 | 5 | 1.7 | 1.6641 | 1.6227 | 1.6098 | 1.91 | 1.56 | 1.67 |  |  | 87
| 1 | 6 | 1.73 | 1.6955 | 1.6489 | 1.586 | 1.88 | 1.64 | 1.6
| Mean |  | 1.7267 | 1.6902 | 1.6491 | 1.6 | 1.9917 | 1.6017 | 1.62
| Median |  | 1.735 | 1.6947 | 1.6488 | 1.6044 | 1.945 | 1.605 | 1.61
| Std Dev. |  | 0.017512 | 0.015751 | 0.035399 | 0.013819 | 0.11583 | 0.028577 | 0.037947
| Rel.Std.Dev. |  | 1.014% | 0.9319% | 2.147% | 0.8637% | 5.816% | 1.784% | 2.342%
| PDM3 |  | 4.778% | 2.566% | 0.06967% | -2.91% | 20.86% | -2.808% | -1.695%
| Z-Score (Absolute) |  | 1.1 | 0.58 | 0 | 0.7 | 4.87 | 0.67 | 0.41
| Analytical results for Mn in OREAS 24b (Certified Value 0.044 wt.%) | [10 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 3
3
| 1 | 1 | 0.0471 | 0.0468 | 0.0433 | 0.0391 | 0.0437 | 0.04285 | 0.0463 | 0.0449 | 0.0438 | 1
| 1 | 2 | 0.0464 | 0.045 | 0.0428 | 0.0398 | 0.0447 | 0.042479 | 0.0491 | 0.0441 | 0.0433 | 0
| 1 | 3 | 0.0474 | 0.0444 | 0.0426 | 0.0419 | 0.0463 | 0.042488 | 0.0466 | 0.0452 | 0.0447 | 16
| 1 | 4 | 0.0472 | 0.0452 | 0.0432 | 0.0414 | 0.0458 | 0.042254 | 0.0458 | 0.0446 | 0.0436 | 0.044226
| 1 | 5 | 0.0462 | 0.0462 | 0.042 | 0.0415 | 0.0441 | 0.042704 | 0.0451 | 0.0433 | 0.0444 | 88
| 1 | 6 | 0.0473 | 0.0466 | 0.0429 | 0.0396 | 0.0448 | 0.042036 | 0.0442 | 0.0454 | 0.0437
| Mean |  | 0.046933 | 0.0457 | 0.0428 | 0.04055 | 0.0449 | 0.042469 | 0.046183 | 0.044583 | 0.043917
| Median |  | 0.04715 | 0.0457 | 0.04285 | 0.0406 | 0.04475 | 0.042483 | 0.04605 | 0.04475 | 0.04375
| Std Dev. |  | 0.00050465 | 0.00096954 | 0.00046904 | 0.0011845 | 0.00098995 | 0.00029488 | 0.0016702 | 0.00077825 | 0.00052694
| Rel.Std.Dev. |  | 1.075% | 2.122% | 1.096% | 2.921% | 2.205% | 0.6943% | 3.617% | 1.746% | 1.2%
| PDM3 |  | 6.121% | 3.333% | -3.225% | -8.312% | 1.524% | -3.974% | 4.425% | 0.8077% | -0.6997%
| Z-Score (Absolute) |  | 0.99 | 0.47 | 0.75 | 1.7 | 0.13 | 0.89 | 0.67 | 0 | 0.28
| Analytical results for Mo in OREAS 24b (Certified Value 4.03 ppm) | [10 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | 3 | 3.5 | 4 | 4.86 | 4 | 4.1904 | 0.8 | 4.03 | 3.7 | 1
| 1 | 2 | 3 | 3.5 | 3.8 | 4.4 | 4.2 | 3.9589 | 0.7 | 3.61 | 3.9 | 9
| 1 | 3 | 4 | 4 | 4 | 4.89 | 4.3 | 4.0433 | 0.7 | 3.66 | 3.8 | 16
| 1 | 4 | 3 | 3.5 | 3.8 | 4.72 | 4.4 | 4.0985 | 0.6 | 3.6 | 3.9 | 4.0324
| 1 | 5 | 3 | 4 | 3.8 | 4.34 | 4.2 | 3.9893 | 0.8 | 3.59 | 3.9 | 89
| 1 | 6 | 3 | 4 | 3.9 | 4.76 | 4.2 | 4.0517 | 1.6 | 3.77 | 4.5
| Mean |  | 3.1667 | 3.75 | 3.8833 | 4.6617 | 4.2167 | 4.0554 | 0.86667 | 3.71 | 3.95
| Median |  | 3 | 3.75 | 3.85 | 4.74 | 4.2 | 4.0475 | 0.75 | 3.635 | 3.9
| Std Dev. |  | 0.40825 | 0.27386 | 0.098319 | 0.23515 | 0.13292 | 0.082363 | 0.36697 | 0.17029 | 0.28107
| Rel.Std.Dev. |  | 12.89% | 7.303% | 2.532% | 5.044% | 3.152% | 2.031% | 42.34% | 4.59% | 7.116%
| PDM3 |  | -21.47% | -7.004% | -3.697% | 15.6% | 4.569% | 0.5686% | -78.51% | -7.996% | -2.044%
| Z-Score (Absolute) |  | 2.79 | 0.52 | 0 | 3.03 | 1.3 | 0.67 | 11.74 | 0.67 | 0.26
| Analytical results for Na in OREAS 24b (Certified Value 0.846 wt.%) | [10 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 3
3
| 1 | 1 | 0.88 | 0.88 | 0.8646 | 0.72 | 0.8083 | 0.82626 | 1.07 | 0.83 | 0.819 | 1
| 1 | 2 | 0.87 | 0.86 | 0.8558 | 0.72 | 0.871 | 0.81523 | 1.25 | 0.81 | 0.808 | 0
| 1 | 3 | 0.88 | 0.84 | 0.859 | 0.75 | 0.8645 | 0.82156 | 1.19 | 0.84 | 0.822 | 16
| 1 | 4 | 0.88 | 0.85 | 0.8649 | 0.74 | 0.8896 | 0.80865 | 1.08 | 0.83 | 0.813 | 0.84574
| 1 | 5 | 0.86 | 0.87 | 0.8678 | 0.73 | 0.8446 | 0.82755 | 1.13 | 0.8 | 0.841 | 90
| 1 | 6 | 0.87 | 0.86 | 0.8644 | 0.71 | 0.8514 | 0.81382 | 1.05 | 0.86 | 0.829
| Mean |  | 0.87333 | 0.86 | 0.86275 | 0.72833 | 0.8549 | 0.81885 | 1.1283 | 0.82833 | 0.822
| Median |  | 0.875 | 0.86 | 0.8645 | 0.725 | 0.85795 | 0.81839 | 1.105 | 0.83 | 0.8205
| Std Dev. |  | 0.008165 | 0.014142 | 0.0044415 | 0.01472 | 0.027732 | 0.0074872 | 0.078081 | 0.02137 | 0.011798
| Rel.Std.Dev. |  | 0.9349% | 1.644% | 0.5148% | 2.021% | 3.244% | 0.9144% | 6.92% | 2.58% | 1.435%
| PDM3 |  | 3.263% | 1.686% | 2.012% | -13.88% | 1.083% | -3.18% | 33.41% | -2.058% | -2.807%
| Z-Score (Absolute) |  | 0.47 | 0.13 | 0.2 | 3.21 | 0 | 0.92 | 6.94 | 0.67 | 0.84
| Analytical results for Nb in OREAS 24b (Certified Value 14.6 ppm) | [10 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 16 | 14.7 | 13.8 | 14.7 | 14.1692 | 0.5 | 15 | 13.1 |  | 1
| 1 | 2 | 17 | 14.5 | 12.9 | 14.3 | 13.8151 | 1.6 | 14.6 | 12.7 |  | 27
| 1 | 3 | 15.5 | 14.8 | 14.1 | 15.3 | 13.9862 | 0.9 | 14.9 | 13 |  | 16
| 1 | 4 | 17 | 14.7 | 14 | 17 | 13.9719 | 1.6 | 14.5 | 12.5 |  | 14.5979
| 1 | 5 | 17.5 | 14.7 | 13.7 | 16 | 13.8429 | 1.8 | 14.2 | 12.8 |  | 91
| 1 | 6 | 17.5 | 14.8 | 14.5 | 15.3 | 13.8259 | 8.6 | 15.1 | 12.8
| Mean |  | 16.75 | 14.7 | 13.8333 | 15.4333 | 13.9352 | 2.5 | 14.7167 | 12.8167
| Median |  | 17 | 14.7 | 13.9 | 15.3 | 13.9074 | 1.6 | 14.75 | 12.8
| Std Dev. |  | 0.82158 | 0.10954 | 0.53541 | 0.96264 | 0.13682 | 3.0292 | 0.34303 | 0.2137
| Rel.Std.Dev. |  | 4.905% | 0.7452% | 3.87% | 6.237% | 0.9818% | 121.2% | 2.331% | 1.667%
| PDM3 |  | 14.74% | 0.6995% | -5.237% | 5.723% | -4.54% | -82.87% | 0.8137% | -12.2%
| Z-Score (Absolute) |  | 2.05 | 0.32 | 0.41 | 0.94 | 0.32 | 9.96 | 0.34 | 1.27
| Analytical results for Nd in OREAS 24b (Indicative Value 36.2 ppm) | [10 empty] | Indicative
Nd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab |  | 1
| No | No | 02: 02; 03: 04; 04: 07 |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*MS |  | 1
1
| 1 | 1 | 02: 35.2; 03: 36.8; 04: 35.4 |  | 1
| 1 | 2 | 02: 35.3; 03: 34; 04: 38.3 |  | 18
| 1 | 3 | 02: 36.9; 03: 38.1; 04: 37.3 |  | 16
| 1 | 4 | 02: 36.2; 03: 37.4; 04: 33.9 |  | 36.2389
| 1 | 5 | 02: 35.4; 03: 38; 04: 36.8 |  | 24
| 1 | 6 | 02: 35.7; 03: 36.7; 04: 34.9
| Mean |  | 02: 35.7833; 03: 36.8333; 04: 36.1
| Median |  | 02: 35.55; 03: 37.1; 04: 36.1
| Std Dev. |  | 02: 0.65549; 03: 1.5055; 04: 1.6456
| Rel.Std.Dev. |  | 02: 1.832%; 03: 4.087%; 04: 4.558%
| PDM3 |  | 02: -1.257%; 03: 1.64%; 04: -0.3833%
| Z-Score (Absolute) |  | 02: 0.67; 03: 1.56; 04: 0
| Analytical results for Ni in OREAS 24b (Certified Value 60 ppm) | [10 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 0
1
| 1 | 1 | 58 | 68 | 65 | 59.9 | 57 | 54.9221 | 66.1 | 59.2 | 66.1 | 1
| 1 | 2 | 57 | 58 | 66 | 58.5 | 58 | 55.0086 | 75.8 | 57.8 | 64.3 | 3
| 1 | 3 | 57 | 60 | 66 | 58.2 | 58 | 54.065 | 67 | 61.1 | 66.7 | 16
| 1 | 4 | 57 | 62 | 66 | 56.6 | 59 | 56.3538 | 71.3 | 59.7 | 67 | 60.0776
| 1 | 5 | 59 | 62 | 66 | 58.3 | 56 | 54.7424 | 66.9 | 56.8 | 69.9 | 92
| 1 | 6 | 58 | 60 | 68 | 59 | 57 | 54.3323 | 67.9 | 60.5 | 64.3
| Mean |  | 57.6667 | 61.6667 | 66.1667 | 58.4167 | 57.5 | 54.9041 | 69.1667 | 59.1833 | 66.3833
| Median |  | 57.5 | 61 | 66 | 58.4 | 57.5 | 54.8323 | 67.45 | 59.45 | 66.4
| Std Dev. |  | 0.8165 | 3.4448 | 0.98319 | 1.087 | 1.0488 | 0.79623 | 3.7254 | 1.629 | 2.0789
| Rel.Std.Dev. |  | 1.416% | 5.586% | 1.486% | 1.861% | 1.824% | 1.45% | 5.386% | 2.752% | 3.132%
| PDM3 |  | -4.013% | 2.645% | 10.14% | -2.765% | -4.29% | -8.611% | 15.13% | -1.489% | 10.5%
| Z-Score (Absolute) |  | 0.41 | 0.67 | 1.9 | 0.21 | 0.46 | 1.16 | 2.71 | 0 | 1.96
| Analytical results for P in OREAS 24b (Certified Value 0.069 wt.%) | [10 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 3
3
| 1 | 1 | 0.076 | 0.08 | 0.0651 | 0.063 | 0.0645 | 0.070156 | 0.072 | 0.066 |  | 1
| 1 | 2 | 0.075 | 0.075 | 0.0653 | 0.062 | 0.0679 | 0.070978 | 0.07 | 0.066 |  | 10
| 1 | 3 | 0.076 | 0.075 | 0.0645 | 0.064 | 0.0683 | 0.070458 | 0.072 | 0.065 |  | 16
| 1 | 4 | 0.075 | 0.075 | 0.0653 | 0.063 | 0.0699 | 0.070169 | 0.071 | 0.065 |  | 0.069287
| 1 | 5 | 0.075 | 0.08 | 0.0649 | 0.063 | 0.0654 | 0.07164 | 0.07 | 0.067 |  | 93
| 1 | 6 | 0.076 | 0.075 | 0.0642 | 0.061 | 0.0657 | 0.069586 | 0.073 | 0.07
| Mean |  | 0.0755 | 0.076667 | 0.064883 | 0.062667 | 0.06695 | 0.070498 | 0.071333 | 0.0665
| Median |  | 0.0755 | 0.075 | 0.065 | 0.063 | 0.0668 | 0.070313 | 0.0715 | 0.066
| Std Dev. |  | 0.00054772 | 0.002582 | 0.00044907 | 0.0010328 | 0.0020686 | 0.00072022 | 0.0012111 | 0.0018708
| Rel.Std.Dev. |  | 0.7255% | 3.368% | 0.6921% | 1.648% | 3.09% | 1.022% | 1.698% | 2.813%
| PDM3 |  | 8.967% | 10.65% | -6.356% | -9.555% | -3.373% | 1.747% | 2.953% | -4.023%
| Z-Score (Absolute) |  | 1.42 | 1.66 | 0.8 | 1.27 | 0.37 | 0.37 | 0.55 | 0.46
| Analytical results for Pb in OREAS 24b (Certified Value 23.1 ppm) | [10 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 23 | 21 | 26 | 24 | 22.0538 | 24.5 | 22.8 | 23.6 |  | 1
| 1 | 2 | 23 | 22 | 25 | 23 | 21.6995 | 24.4 | 22.1 | 23 |  | 11
| 1 | 3 | 23 | 24 | 26 | 25 | 23.5 | 26.4 | 22.5 | 23.2 |  | 16
| 1 | 4 | 23 | 21 | 25 | 24 | 22.1347 | 23.3 | 22.3 | 21.8 |  | 23.1472
| 1 | 5 | 21 | 21 | 26 | 23 | 22.1698 | 24.6 | 22.2 | 23.5 |  | 94
| 1 | 6 | 23 | 22 | 24 | 23 | 22.1095 | 24 | 22.7 | 22.1
| Mean |  | 22.6667 | 21.8333 | 25.3333 | 23.6667 | 22.2779 | 24.5333 | 22.4333 | 22.8667
| Median |  | 23 | 21.5 | 25.5 | 23.5 | 22.1221 | 24.45 | 22.4 | 23.1
| Std Dev. |  | 0.8165 | 1.169 | 0.8165 | 0.8165 | 0.62272 | 1.0309 | 0.28048 | 0.74744
| Rel.Std.Dev. |  | 3.602% | 5.354% | 3.223% | 3.45% | 2.795% | 4.202% | 1.25% | 3.269%
| PDM3 |  | -2.076% | -5.676% | 9.444% | 2.244% | -3.756% | 5.988% | -3.084% | -1.212%
| Z-Score (Absolute) |  | 0.1 | 0.91 | 2.49 | 0.87 | 0.47 | 1.72 | 0.32 | 0.1
| Analytical results for Pr in OREAS 24b (Indicative Value 9.86 ppm) | [10 empty] | Indicative
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 07 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS |  | 2
2
| 1 | 1 | 01: 9.58; 02: 9.42; 03: 10.2 |  | 1
| 1 | 2 | 01: 9.8; 02: 9.17; 03: 11 |  | 19
| 1 | 3 | 01: 9.76; 02: 9.63; 03: 10.6 |  | 16
| 1 | 4 | 01: 9.58; 02: 9.77; 03: 9.9 |  | 9.86
| 1 | 5 | 01: 9.56; 02: 9.48; 03: 10.8 |  | 25
| 1 | 6 | 01: 9.6; 02: 9.73; 03: 9.9
| Mean |  | 01: 9.6467; 02: 9.5333; 03: 10.4
| Median |  | 01: 9.59; 02: 9.555; 03: 10.4
| Std Dev. |  | 01: 0.10482; 02: 0.22438; 03: 0.46904
| Rel.Std.Dev. |  | 01: 1.087%; 02: 2.354%; 03: 4.51%
| PDM3 |  | 01: -2.164%; 02: -3.313%; 03: 5.477%
| Z-Score (Absolute) |  | 01: 0; 02: 0.67; 03: 4.48
| Analytical results for Rb in OREAS 24b (Certified Value 164 ppm) | [10 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 158 | 162.3 | 158 | 172.6 | 159.8663 | 175 | 169 | 162.8 |  | 1
| 1 | 2 | 160 | 158.8 | 153 | 172.5 | 157.0843 | 203 | 168 | 158.8 |  | 0
| 1 | 3 | 162 | 155 | 148 | 176.7 | 159.1044 | 176 | 168.5 | 162.1 |  | 16
| 1 | 4 | 167 | 152.6 | 155 | 175.1 | 158.7296 | 184 | 170.5 | 156.4 |  | 163.6294
| 1 | 5 | 165 | 147.4 | 145 | 173.1 | 158.4771 | 129 | 161.5 | 165.8 |  | 95
| 1 | 6 | 162 | 159.1 | 158 | 167.1 | 157.1502 | 174 | 174.5 | 159.1
| Mean |  | 162.3333 | 155.8667 | 152.8333 | 172.85 | 158.402 | 173.5 | 168.6667 | 160.8333
| Median |  | 162 | 156.9 | 154 | 172.85 | 158.6034 | 175.5 | 168.75 | 160.6
| Std Dev. |  | 3.266 | 5.3582 | 5.3448 | 3.2617 | 1.1004 | 24.3701 | 4.2269 | 3.3744
| Rel.Std.Dev. |  | 2.012% | 3.438% | 3.497% | 1.887% | 0.6947% | 14.05% | 2.506% | 2.098%
| PDM3 |  | -0.7921% | -4.744% | -6.598% | 5.635% | -3.195% | 6.032% | 3.078% | -1.709%
| Z-Score (Absolute) |  | 0.08 | 0.6 | 0.92 | 1.19 | 0.34 | 1.26 | 0.75 | 0.08
| Analytical results for Re in OREAS 24b (Indicative Value 0.002 ppm) | [10 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 05 | 07 | 08 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS |  |  |  |  | 3
3
| 1 | 1 | < 0.1 | 0.7 | < 0.05 | 0.003 | < 0.002 |  |  |  |  | 1
| 1 | 2 | < 0.1 | 0.9 | < 0.05 | < 0.001 | < 0.002 |  |  |  |  | 12
| 1 | 3 | < 0.1 | 0.7 | < 0.05 | 0.003 | < 0.002 |  |  |  |  | 16
| 1 | 4 | < 0.1 | 0.6 | < 0.05 | < 0.001 | < 0.002 |  |  |  |  | 0.0015833
| 1 | 5 | < 0.1 | 0.7 | < 0.05 | 0.004 | 0.002 |  |  |  |  | 26
| 1 | 6 | < 0.1 | 0.7 | < 0.05 | 0.002 | < 0.002
| Mean |  | 03: 0.71667; 05: 0.003; 06: 0.002
| Median |  | 03: 0.7; 05: 0.003; 06: 0.002
| Std Dev. |  | 03: 0.098319; 05: 0.0008165
| Rel.Std.Dev. |  | 03: 13.72%; 05: 27.22%
| PDM3 |  | 03: 45163%; 05: 89.47%; 06: 26.32%
| Z-Score (Absolute) |  | 0.71 | 19.57 | 0 | 0.65 | 0.67
| Analytical results for S in OREAS 24b (Certified Value 0.198 wt.%) | [10 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 3
3
| 1 | 1 | 0.21 | 0.205 | 0.2041 | 0.24 | 0.1868 | 0.18777 | 0.2 | 0.2 |  | 1
| 1 | 2 | 0.21 | 0.195 | 0.1976 | 0.22 | 0.1875 | 0.18531 | 0.2 | 0.2 |  | 4
| 1 | 3 | 0.21 | 0.2 | 0.1986 | 0.22 | 0.1837 | 0.18847 | 0.2 | 0.2 |  | 16
| 1 | 4 | 0.21 | 0.2 | 0.2013 | 0.23 | 0.1967 | 0.18888 | 0.2 | 0.2 |  | 0.1983
| 1 | 5 | 0.21 | 0.21 | 0.1958 | 0.23 | 0.191 | 0.19088 | 0.2 | 0.2 |  | 96
| 1 | 6 | 0.21 | 0.205 | 0.2004 | 0.22 | 0.1861 | 0.18253 | 0.21 | 0.2
| Mean |  | 0.21 | 0.2025 | 0.19963 | 0.22667 | 0.18863 | 0.18731 | 0.20167 | 0.2
| Median |  | 0.21 | 0.2025 | 0.1995 | 0.225 | 0.18715 | 0.18812 | 0.2 | 0.2
| Std Dev. |  | 0 | 0.005244 | 0.0029398 | 0.008165 | 0.0046059 | 0.0029548 | 0.0040825 | 3.0405e-17
| Rel.Std.Dev. |  | 0% | 2.59% | 1.473% | 3.602% | 2.442% | 1.578% | 2.024% | 1.52e-14%
| PDM3 |  | 5.898% | 2.116% | 0.6704% | 14.3% | -4.877% | -5.546% | 1.696% | 0.8553%
| Z-Score (Absolute) |  | 1.14 | 0.21 | 0.15 | 3.22 | 1.52 | 1.68 | 0.1 | 0.1
| Analytical results for Sb in OREAS 24b (Certified Value 1 ppm) | [10 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | < 5 | 0.8 | 1 | 1.38 | 1 | 1.0356 | < 0.1 | 0.99 | 1.1 | 1
| 1 | 2 | < 5 | 0.8 | 1 | 1.17 | 1.1 | 0.98595 | < 0.1 | 0.98 | 0.9 | 7
| 1 | 3 | < 5 | 0.8 | 1 | 1.07 | 1 | 0.9803 | < 0.1 | 0.98 | 1 | 16
| 1 | 4 | < 5 | 0.8 | 1 | 1.5 | 1.1 | 1.005 | < 0.1 | 0.96 | 1 | 0.99725
| 1 | 5 | < 5 | 1 | 1 | 1.02 | 1 | 0.99024 | < 0.1 | 0.95 | 1 | 97
| 1 | 6 | < 5 | 1 | 1 | 1.21 | 1 | 0.97811 | 0.3 | 1.03 | 0.9
| Mean |  |  | 0.86667 | 1 | 1.225 | 1.0333 | 0.99587 | 0.3 | 0.98167 | 0.98333
| Median |  |  | 0.8 | 1 | 1.19 | 1 | 0.98809 | 0.3 | 0.98 | 1
| Std Dev. |  |  | 0.10328 | 0 | 0.18382 | 0.05164 | 0.021673 |  | 0.027869 | 0.075277
| Rel.Std.Dev. |  |  | 11.92% | 0% | 15.01% | 4.997% | 2.176% |  | 2.839% | 7.655%
| PDM3 |  |  | -13.09% | 0.2756% | 22.84% | 3.618% | -0.1389% | -69.92% | -1.563% | -1.396%
| Z-Score (Absolute) |  | 27.07 | 2.33 | 0.07 | 4.12 | 0.67 | 0 | 16.27 | 0.26 | 0.23
| Analytical results for Sc in OREAS 24b (Certified Value 15.3 ppm) | [10 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 09 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  | 1
2
| 1 | 1 | 16 | 16 | 14 | 16.5 | 14 | 18.8087 | 15.8 | 13 |  | 1
| 1 | 2 | 16 | 15 | 14 | 15.6 | 15 | 18.3514 | 15.5 | 13 |  | 0
| 1 | 3 | 16 | 15 | 14 | 16.3 | 15 | 17.8473 | 15.8 | 13 |  | 16
| 1 | 4 | 16 | 16 | 14 | 16.8 | 15 | 17.0191 | 15.6 | 13 |  | 15.2548
| 1 | 5 | 15 | 16 | 13 | 16.6 | 14 | 16.2507 | 15.1 | 13 |  | 98
| 1 | 6 | 16 | 15 | 14 | 16.6 | 14 | 15.7522 | 16 | 13
| Mean |  | 15.8333 | 15.5 | 13.8333 | 16.4 | 14.5 | 17.3382 | 15.6333 | 13
| Median |  | 16 | 15.5 | 14 | 16.55 | 14.5 | 17.4332 | 15.7 | 13
| Std Dev. |  | 0.40825 | 0.54772 | 0.40825 | 0.42426 | 0.54772 | 1.204 | 0.31411 | 0
| Rel.Std.Dev. |  | 2.578% | 3.534% | 2.951% | 2.587% | 3.777% | 6.944% | 2.009% | 0%
| PDM3 |  | 3.793% | 1.608% | -9.318% | 7.507% | -4.948% | 13.66% | 2.482% | -14.78%
| Z-Score (Absolute) |  | 0.19 | 0.05 | 1.23 | 0.59 | 0.76 | 1.26 | 0.05 | 1.82
| Analytical results for Se in OREAS 24b (Indicative Value 0.66 ppm) | [10 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
2
| 1 | 1 | < 5 | 0.5 | < 2 | < 1 | 0.62115 | 0.7 | 1 |  |  | 1
| 1 | 2 | < 5 | < 0.5 | < 2 | < 1 | 0.25514 | 1.2 | 1 |  |  | 8
| 1 | 3 | < 5 | < 0.5 | < 2 | < 1 | 0.37931 | 0.7 | 1 |  |  | 16
| 1 | 4 | < 5 | 0.5 | < 2 | < 1 | 0.47337 | 0.6 | 1 |  |  | 0.66185
| 1 | 5 | < 5 | < 0.5 | < 2 | < 1 | 0.32878 | 1.1 | 1 |  |  | 20
| 1 | 6 | < 5 | 0.5 | < 2 | < 1 | 0.27662 | 1 | 1
| Mean |  | 03: 0.5; 06: 0.38906; 07: 0.88333; 08: 1
| Median |  | 03: 0.5; 06: 0.35405; 07: 0.85; 08: 1
| Std Dev. |  | 03: 0; 06: 0.13805; 07: 0.24833; 08: 0
| Rel.Std.Dev. |  | 03: 0%; 06: 35.48%; 07: 28.11%; 08: 0%
| PDM3 |  | 03: -24.45%; 06: -41.22%; 07: 33.46%; 08: 51.09%
| Z-Score (Absolute) |  | 2.84 | 0.89 | 0.21 | 0.67 | 0.87 | 0 | 0.21
| Analytical results for Sm in OREAS 24b (Indicative Value 7.06 ppm) | [10 empty] | Indicative
Sm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab |  |  | 1
| No | No | 02: 02; 03: 04; 04: 07 |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*MS |  |  | 2
2
| 1 | 1 | 02: 7.55; 03: 6.9; 04: 6.9 |  |  | 1
| 1 | 2 | 02: 7.55; 03: 6.4; 04: 7.4 |  |  | 15
| 1 | 3 | 02: 7.7; 03: 6.7; 04: 7.1 |  |  | 16
| 1 | 4 | 02: 7.7; 03: 6.9; 04: 6.6 |  |  | 7.0583
| 1 | 5 | 02: 7.3; 03: 6.7; 04: 6.9 |  |  | 21
| 1 | 6 | 02: 7.55; 03: 6.4; 04: 6.8
| Mean |  | 02: 7.5583; 03: 6.6667; 04: 6.95
| Median |  | 02: 7.55; 03: 6.7; 04: 6.9
| Std Dev. |  | 02: 0.14634; 03: 0.22509; 04: 0.27386
| Rel.Std.Dev. |  | 02: 1.936%; 03: 3.376%; 04: 3.94%
| PDM3 |  | 02: 7.084%; 03: -5.549%; 04: -1.535%
| Z-Score (Absolute) |  | 02: 1.45; 03: 0.67; 04: 0
| Analytical results for Sn in OREAS 24b (Certified Value 4.25 ppm) | [10 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
3
| 1 | 1 | 4 | 4.6 | 4.7 | 4.4 | 4.19 | 3 | 4.4 | 4.1 |  | 1
| 1 | 2 | 4 | 4.5 | 4.3 | 4.5 | 4.54 | 3 | 4.3 | 4.2 |  | 0
| 1 | 3 | 4 | 4.6 | 4.2 | 4.4 | 4.14 | 3 | 4.4 | 4 |  | 16
| 1 | 4 | 4 | 4.5 | 4.3 | 4.5 | 4.09 | 5 | 4.2 | 3.8 |  | 4.2509
| 1 | 5 | 4 | 4.5 | 4.2 | 4.5 | 4.04 | 4 | 4.3 | 4.1 |  | 99
| 1 | 6 | 4 | 4.5 | 4.3 | 4.3 | 4.07 | 3 | 4.4 | 3.9
| Mean |  | 4 | 4.5333 | 4.3333 | 4.4333 | 4.1783 | 3.5 | 4.3333 | 4.0167
| Median |  | 4 | 4.5 | 4.3 | 4.45 | 4.115 | 3 | 4.35 | 4.05
| Std Dev. |  | 0 | 0.05164 | 0.18619 | 0.08165 | 0.18498 | 0.83666 | 0.08165 | 0.1472
| Rel.Std.Dev. |  | 0% | 1.139% | 4.297% | 1.842% | 4.427% | 23.9% | 1.884% | 3.665%
| PDM3 |  | -5.901% | 6.645% | 1.94% | 4.293% | -1.706% | -17.66% | 1.94% | -5.509%
| Z-Score (Absolute) |  | 0.83 | 0.9 | 0.25 | 0.57 | 0.25 | 2.45 | 0.25 | 0.77
| Analytical results for Sr in OREAS 24b (Certified Value 124 ppm) | [10 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 138 | 117 | 125.1 | 96.8 | 128.3 | 118.9135 | 118 | 126 | 145 | 1
| 1 | 2 | 137 | 119 | 124.9 | 101 | 129.4 | 117.6757 | 140 | 124 | 138 | 5
| 1 | 3 | 138 | 122 | 126 | 103 | 134.8 | 118.0493 | 121 | 126 | 139 | 16
| 1 | 4 | 138 | 122 | 124.9 | 103 | 136 | 116.2714 | 123 | 123.5 | 133 | 123.8498
| 1 | 5 | 135 | 119 | 123.7 | 106 | 132.2 | 117.9317 | 109 | 121 | 141 | 100
| 1 | 6 | 136 | 117 | 127.4 | 101 | 130.7 | 116.2488 | 117 | 128.5 | 140
| Mean |  | 137 | 119.3333 | 125.3333 | 101.8 | 131.9 | 117.515 | 121.3333 | 124.8333 | 139.3333
| Median |  | 137.5 | 119 | 125 | 102 | 131.45 | 117.8037 | 119.5 | 125 | 139.5
| Std Dev. |  | 1.2649 | 2.2509 | 1.2501 | 3.0594 | 3.0318 | 1.0574 | 10.328 | 2.582 | 3.9328
| Rel.Std.Dev. |  | 0.9233% | 1.886% | 0.9974% | 3.005% | 2.299% | 0.8998% | 8.512% | 2.068% | 2.823%
| PDM3 |  | 10.62% | -3.647% | 1.198% | -17.8% | 6.5% | -5.115% | -2.032% | 0.7941% | 12.5%
| Z-Score (Absolute) |  | 1.16 | 0.52 | 0.05 | 2.2 | 0.67 | 0.7 | 0.33 | 0 | 1.38
| Analytical results for Ta in OREAS 24b (Certified Value 1.23 ppm) | [10 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 2
2
| 1 | 1 | 1 | 1.2 | 1.57 | 1.5 | 1.24 | < 0.1 | 1.19 | 1 |  | 1
| 1 | 2 | 1.1 | 1.1 | 1.51 | 1.41 | 1.218 | < 0.1 | 1.16 | 0.9 |  | 26
| 1 | 3 | 1 | 1.1 | 1.63 | 1.46 | 1.226 | < 0.1 | 1.18 | 0.9 |  | 16
| 1 | 4 | 1.1 | 1.1 | 1.4 | 1.74 | 1.234 | < 0.1 | 1.18 | 0.9 |  | 1.2273
| 1 | 5 | 1.1 | 1.1 | 1.61 | 1.54 | 1.227 | < 0.1 | 1.14 | 1 |  | 101
| 1 | 6 | 1.2 | 1 | 1.51 | 1.39 | 1.251 | 0.2 | 1.21 | 0.9
| Mean |  | 1.0833 | 1.1 | 1.5383 | 1.5067 | 1.2327 | 0.2 | 1.1767 | 0.93333
| Median |  | 1.1 | 1.1 | 1.54 | 1.48 | 1.2305 | 0.2 | 1.18 | 0.9
| Std Dev. |  | 0.075277 | 0.063246 | 0.084004 | 0.12707 | 0.01169 |  | 0.024221 | 0.05164
| Rel.Std.Dev. |  | 6.949% | 5.75% | 5.461% | 8.434% | 0.9484% |  | 2.058% | 5.533%
| PDM3 |  | -11.73% | -10.37% | 25.34% | 22.76% | 0.4384% | -83.7% | -4.124% | -23.95%
| Z-Score (Absolute) |  | 0.25 | 0.17 | 1.8 | 1.66 | 0.43 | 4.79 | 0.17 | 0.92
| Analytical results for Tb in OREAS 24b (Indicative Value 0.87 ppm) | [10 empty] | Indicative
Tb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 02: 02; 03: 04; 04: 06; 05: 07 |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS |  | 2
2
| 1 | 1 | 02: 0.84; 03: 0.73; 04: 0.84615; 05: 0.9 |  | 1
| 1 | 2 | 02: 0.9; 03: 0.67; 04: 0.83351; 05: 0.9 |  | 16
| 1 | 3 | 02: 0.86; 03: 0.62; 04: 0.86798; 05: 0.9 |  | 16
| 1 | 4 | 02: 0.84; 03: 0.62; 04: 0.8593; 05: 0.9 |  | 0.8655
| 1 | 5 | 02: 0.82; 03: 0.64; 04: 0.86341; 05: 0.8 |  | 22
| 1 | 6 | 02: 0.88; 03: 0.62; 04: 0.86866; 05: 0.9
| Mean |  | 02: 0.85667; 03: 0.65; 04: 0.8565; 05: 0.88333
| Median |  | 02: 0.85; 03: 0.63; 04: 0.86136; 05: 0.9
| Std Dev. |  | 02: 0.029439; 03: 0.043818; 04: 0.013933; 05: 0.040825
| Rel.Std.Dev. |  | 02: 3.436%; 03: 6.741%; 04: 1.627%; 05: 4.622%
| PDM3 |  | 02: -1.021%; 03: -24.9%; 04: -1.04%; 05: 2.06%
| Z-Score (Absolute) |  | 02: 0; 03: 10.38; 04: 0; 05: 1.34
| Analytical results for Te in OREAS 24b (Indicative Value < 0.1 ppm) | [10 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | < 0.2 | 0.2 | < 0.1 | 0.024038 | < 0.1 | < 0.05 |  |  |  | 1
| 1 | 2 | < 0.2 | 0.2 | < 0.1 | 0.027027 | < 0.1 | 0.05 |  |  |  | 9
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 0.030542 | < 0.1 | < 0.05 |  |  |  | 16
| 1 | 4 | < 0.2 | 0.2 | < 0.1 | 0.0060302 | < 0.1 | < 0.05 |  |  |  | < 0.1
| 1 | 5 | < 0.2 | < 0.2 | < 0.1 | 0.02439 | < 0.1 | < 0.05 |  |  |  | 23
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 0.025871 | < 0.1 | < 0.05
| Mean |  | 03: 0.2; 06: 0.022983; 08: 0.05
| Median |  | 03: 0.2; 06: 0.02513; 08: 0.05
| Std Dev. |  | 03: 3.3993e-17; 06: 0.0086293
| Rel.Std.Dev. |  | 03: 1.7e-14%; 06: 37.55%
PDM3
| Z-Score (Absolute) |  | 1.41 | 2.82 | 0 | 0.76 | 0 | 0.59
| Analytical results for Th in OREAS 24b (Certified Value 16.4 ppm) | [10 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 1
2
| 1 | 1 | 20 | 15.7 | 14.6 | 17.4 | 16.16 | 16.89 | 18.4 | 18 | 17 | 1
| 1 | 2 | 20 | 16 | 14.9 | 16.2 | 15.99 | 16.76 | 17.2 | 17.8 | 16.4 | 0
| 1 | 3 | 20 | 16.3 | 14.7 | 16.1 | 16.49 | 17.18 | 18.2 | 17.8 | 16.3 | 16
| 1 | 4 | 20 | 15.1 | 14.5 | 15.8 | 16.25 | 17.08 | 17.1 | 17.8 | 15.5 | 16.421
| 1 | 5 | 20 | 14.8 | 14.5 | 16.2 | 16.56 | 17.2 | 17.1 | 17.3 | 16.4 | 102
| 1 | 6 | 20 | 15.2 | 15.2 | 16.3 | 15.82 | 17.11 | 18.8 | 18.1 | 15.3
| Mean |  | 20 | 15.5167 | 14.7333 | 16.3333 | 16.2117 | 17.0367 | 17.8 | 17.8 | 16.15
| Median |  | 20 | 15.45 | 14.65 | 16.2 | 16.205 | 17.095 | 17.7 | 17.8 | 16.35
| Std Dev. |  | 0 | 0.57764 | 0.27325 | 0.55015 | 0.28477 | 0.17466 | 0.75631 | 0.27568 | 0.63482
| Rel.Std.Dev. |  | 0% | 3.723% | 1.855% | 3.368% | 1.757% | 1.025% | 4.249% | 1.549% | 3.931%
| PDM3 |  | 21.79% | -5.507% | -10.28% | -0.5341% | -1.275% | 3.749% | 8.398% | 8.398% | -1.651%
| Z-Score (Absolute) |  | 3.03 | 0.67 | 1.32 | 0 | 0.1 | 0.58 | 1.21 | 1.21 | 0.15
| Analytical results for Ti in OREAS 24b (Certified Value 0.468 wt.%) | [10 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 08 | 09 |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.48 | 0.4792 | 0.4273 | 0.41545 | 0.453 | 0.499 |  |  |  | 1
| 1 | 2 | 0.47 | 0.477 | 0.4529 | 0.41965 | 0.452 | 0.495 |  |  |  | 0
| 1 | 3 | 0.47 | 0.4728 | 0.4616 | 0.41438 | 0.455 | 0.498 |  |  |  | 16
| 1 | 4 | 0.47 | 0.4755 | 0.5017 | 0.40723 | 0.455 | 0.49 |  |  |  | 0.46836
| 1 | 5 | 0.46 | 0.4777 | 0.4583 | 0.40947 | 0.438 | 0.498 |  |  |  | 103
| 1 | 6 | 0.47 | 0.4742 | 0.456 | 0.4084 | 0.457 | 0.477
| Mean |  | 0.47 | 0.47607 | 0.45963 | 0.41243 | 0.45167 | 0.49283
| Median |  | 0.47 | 0.47625 | 0.45715 | 0.41193 | 0.454 | 0.4965
| Std Dev. |  | 0.0063246 | 0.0023594 | 0.023997 | 0.0048419 | 0.0069186 | 0.0084242
| Rel.Std.Dev. |  | 1.346% | 0.4956% | 5.221% | 1.174% | 1.532% | 1.709%
| PDM3 |  | 0.3507% | 1.646% | -1.863% | -11.94% | -3.564% | 5.226%
| Z-Score (Absolute) |  | 0.29 | 0.62 | 0.29 | 2.9 | 0.73 | 1.55
| Analytical results for Tl in OREAS 24b (Certified Value 0.86 ppm) | [10 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  | 2
3
| 1 | 1 | < 10 | 0.9 | 0.9 | 0.88 | 0.88942 | 1 | 0.87 |  |  | 1
| 1 | 2 | < 10 | 0.9 | 0.8 | 0.88 | 0.8627 | 1.02 | 0.84 |  |  | 0
| 1 | 3 | < 10 | 0.9 | 0.8 | 0.94 | 0.8867 | 1.06 | 0.86 |  |  | 16
| 1 | 4 | < 10 | 0.9 | 0.8 | 0.9 | 0.89146 | 0.95 | 0.85 |  |  | 0.86207
| 1 | 5 | < 10 | 0.8 | 0.8 | 0.91 | 0.90439 | 0.99 | 0.84 |  |  | 104
| 1 | 6 | < 10 | 0.8 | 0.8 | 0.88 | 0.90746 | 0.97 | 0.87
| Mean |  |  | 0.86667 | 0.81667 | 0.89833 | 0.89036 | 0.99833 | 0.855
| Median |  |  | 0.9 | 0.8 | 0.89 | 0.89044 | 0.995 | 0.855
| Std Dev. |  |  | 0.05164 | 0.040825 | 0.024014 | 0.015936 | 0.038687 | 0.013784
| Rel.Std.Dev. |  |  | 5.958% | 4.999% | 2.673% | 1.79% | 3.875% | 1.612%
| PDM3 |  |  | 0.5331% | -5.267% | 4.206% | 3.281% | 15.81% | -0.8203%
| Z-Score (Absolute) |  | 78.38 | 0.45 | 1.41 | 0.15 | 0 | 2.06 | 0.67
| Analytical results for Tm in OREAS 24b (Indicative Value 0.31 ppm) | [10 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  |  | 1
| No | No | 01: 02; 02: 04; 03: 07 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS |  |  | 2
2
| 1 | 1 | 01: 0.36; 02: 0.17; 03: 0.4 |  |  | 1
| 1 | 2 | 01: 0.4; 02: 0.17; 03: 0.4 |  |  | 18
| 1 | 3 | 01: 0.4; 02: 0.13; 03: 0.4 |  |  | 16
| 1 | 4 | 01: 0.38; 02: 0.14; 03: 0.4 |  |  | 0.30667
| 1 | 5 | 01: 0.36; 02: 0.12; 03: 0.3 |  |  | 24
| 1 | 6 | 01: 0.36; 02: 0.13; 03: 0.4
| Mean |  | 01: 0.37667; 02: 0.14333; 03: 0.38333
| Median |  | 01: 0.37; 02: 0.135; 03: 0.4
| Std Dev. |  | 01: 0.019664; 02: 0.021602; 03: 0.040825
| Rel.Std.Dev. |  | 01: 5.22%; 02: 15.07%; 03: 10.65%
| PDM3 |  | 01: 22.83%; 02: -53.26%; 03: 25%
| Z-Score (Absolute) |  | 01: 0; 02: 23.6; 03: 0.67
| Analytical results for U in OREAS 24b (Certified Value 3.06 ppm) | [10 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
3
| 1 | 1 | < 10 | 3 | 3.1 | 3.7 | 2.99 | 3.16 | 3.1 | 3.1 | 3 | 1
| 1 | 2 | < 10 | 3.1 | 3.1 | 3.6 | 2.96 | 3.17 | 3.1 | 3 | 3 | 0
| 1 | 3 | < 10 | 3.2 | 3.1 | 3.2 | 3 | 3.23 | 3.1 | 3.1 | 2.9 | 16
| 1 | 4 | < 10 | 3 | 3.1 | 3.3 | 2.98 | 3.22 | 3 | 3 | 3 | 3.0593
| 1 | 5 | < 10 | 2.9 | 3.1 | 3.3 | 3.11 | 3.2 | 3 | 3 | 2.9 | 105
| 1 | 6 | < 10 | 3 | 3.2 | 3.3 | 2.96 | 3.25 | 3 | 3.1 | 2.7
| Mean |  |  | 3.0333 | 3.1167 | 3.4 | 3 | 3.205 | 3.05 | 3.05 | 2.9167
| Median |  |  | 3 | 3.1 | 3.3 | 2.985 | 3.21 | 3.05 | 3.05 | 2.95
| Std Dev. |  |  | 0.10328 | 0.040825 | 0.2 | 0.056214 | 0.035071 | 0.054772 | 0.054772 | 0.1169
| Rel.Std.Dev. |  |  | 3.405% | 1.31% | 5.882% | 1.874% | 1.094% | 1.796% | 1.796% | 4.008%
| PDM3 |  |  | -0.8483% | 1.876% | 11.14% | -1.938% | 4.763% | -0.3035% | -0.3035% | -4.662%
| Z-Score (Absolute) |  | 19.72 | 0.17 | 0.67 | 3.54 | 0.51 | 1.57 | 0 | 0 | 1.35
| Analytical results for V in OREAS 24b (Certified Value 108 ppm) | [10 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 116 | 115 | 112 | 80 | 100 | 102.9519 | 78 | 109 | 108 | 1
| 1 | 2 | 113 | 110 | 108 | 80 | 104 | 103.0119 | 82 | 105 | 108 | 0
| 1 | 3 | 114 | 110 | 109 | 81 | 106 | 104.1478 | 75 | 109 | 108 | 16
| 1 | 4 | 115 | 110 | 111 | 79 | 106 | 100.5628 | 61 | 106 | 106 | 108.0611
| 1 | 5 | 112 | 115 | 110 | 80 | 101 | 101.8439 | 83 | 103 | 109 | 106
| 1 | 6 | 114 | 110 | 111 | 79 | 103 | 102.0498 | 91 | 110 | 108
| Mean |  | 114 | 111.6667 | 110.1667 | 79.8333 | 103.3333 | 102.428 | 78.3333 | 107 | 107.8333
| Median |  | 114 | 110 | 110.5 | 80 | 103.5 | 102.5008 | 80 | 107.5 | 108
| Std Dev. |  | 1.4142 | 2.582 | 1.472 | 0.75277 | 2.5033 | 1.2277 | 10.0731 | 2.7568 | 0.98319
| Rel.Std.Dev. |  | 1.241% | 2.312% | 1.336% | 0.9429% | 2.423% | 1.199% | 12.86% | 2.576% | 0.9118%
| PDM3 |  | 5.496% | 3.337% | 1.948% | -26.12% | -4.375% | -5.213% | -27.51% | -0.982% | -0.2108%
| Z-Score (Absolute) |  | 1.03 | 0.69 | 0.47 | 4.01 | 0.54 | 0.67 | 4.23 | 0 | 0.12
| Analytical results for W in OREAS 24b (Certified Value 3.64 ppm) | [10 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 2
2
| 1 | 1 | 10 | 3 | 4.3 | 4.2 | 3.6 | 3.2663 | < 0.1 | 3.9 | 3.1 | 1
| 1 | 2 | < 10 | 4 | 3.6 | 4.2 | 4 | 3.3373 | < 0.1 | 3.7 | 3.2 | 30
| 1 | 3 | < 10 | 3 | 4 | 4.3 | 4.2 | 5.0956 | < 0.1 | 4 | 2.5 | 16
| 1 | 4 | < 10 | 3.5 | 3.8 | 3.7 | 3.7 | 3.595 | < 0.1 | 3.6 | 2.7 | 3.6391
| 1 | 5 | < 10 | 4 | 3.7 | 4 | 3.7 | 3.3541 | 0.2 | 3.6 | 3 | 107
| 1 | 6 | < 10 | 4.5 | 3.8 | 4.1 | 3.6 | 3.6507 | 0.6 | 3.8 | 2.6
| Mean |  | 10 | 3.6667 | 3.8667 | 4.0833 | 3.8 | 3.7165 | 0.4 | 3.7667 | 2.85
| Median |  | 10 | 3.75 | 3.8 | 4.15 | 3.7 | 3.4746 | 0.4 | 3.75 | 2.85
| Std Dev. |  |  | 0.60553 | 0.25033 | 0.2137 | 0.24495 | 0.69263 | 0.28284 | 0.1633 | 0.2881
| Rel.Std.Dev. |  |  | 16.51% | 6.474% | 5.233% | 6.446% | 18.64% | 70.71% | 4.335% | 10.11%
| PDM3 |  | 174.8% | 0.7562% | 6.252% | 12.21% | 4.42% | 2.126% | -89.01% | 3.504% | -21.68%
| Z-Score (Absolute) |  | 13.94 | 0.67 | 0.67 | 2.14 | 0.22 | 0.34 | 24.28 | 0 | 6.18
| Analytical results for Y in OREAS 24b (Certified Value 19.9 ppm) | [10 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 1
1
| 1 | 1 | 22.5 | 20.2 | 16.6 | 20 | 18.25 | 24.6 | 20.3 | 20 |  | 1
| 1 | 2 | 24.6 | 20.5 | 17 | 19.7 | 18.17 | 27.9 | 19.9 | 19.1 |  | 0
| 1 | 3 | 23.7 | 20.5 | 16.2 | 20.6 | 18.76 | 25.2 | 20.2 | 19.4 |  | 16
| 1 | 4 | 23.1 | 19.8 | 17.1 | 21.5 | 18.12 | 24 | 21.1 | 18.5 |  | 19.8712
| 1 | 5 | 24.9 | 19.8 | 17 | 21.1 | 18.53 | 27.1 | 19.3 | 19.8 |  | 108
| 1 | 6 | 22.6 | 20.5 | 16.8 | 20.5 | 18.26 | 24 | 20.5 | 19.6
| Mean |  | 23.5667 | 20.2167 | 16.7833 | 20.5667 | 18.3483 | 25.4667 | 20.2167 | 19.4
| Median |  | 23.4 | 20.35 | 16.9 | 20.55 | 18.255 | 24.9 | 20.25 | 19.5
| Std Dev. |  | 1.0152 | 0.34303 | 0.33714 | 0.66833 | 0.24653 | 1.6561 | 0.60139 | 0.54037
| Rel.Std.Dev. |  | 4.308% | 1.697% | 2.009% | 3.25% | 1.344% | 6.503% | 2.975% | 2.785%
| PDM3 |  | 18.6% | 1.739% | -15.54% | 3.5% | -7.664% | 28.16% | 1.739% | -2.371%
| Z-Score (Absolute) |  | 1.68 | 0 | 1.72 | 0.18 | 0.94 | 2.64 | 0 | 0.41
| Analytical results for Yb in OREAS 24b (Indicative Value 2.17 ppm) | [10 empty] | Indicative
Yb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 02: 02; 03: 04; 04: 06; 05: 07 |  | ppm
|  |  | 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS |  | 2
2
| 1 | 1 | 02: 2.55; 03: 1.9; 04: 1.9865; 05: 2.2 |  | 1
| 1 | 2 | 02: 2.45; 03: 2.1; 04: 1.9924; 05: 2.4 |  | 19
| 1 | 3 | 02: 2.65; 03: 1.9; 04: 2.066; 05: 2.6 |  | 16
| 1 | 4 | 02: 2.35; 03: 1.9; 04: 2.009; 05: 2.2 |  | 2.1659
| 1 | 5 | 02: 2.6; 03: 1.7; 04: 2.0663; 05: 2.1 |  | 25
| 1 | 6 | 02: 2.4; 03: 2; 04: 2.0418; 05: 2.2
| Mean |  | 02: 2.5; 03: 1.9167; 04: 2.027; 05: 2.2833
| Median |  | 02: 2.5; 03: 1.9; 04: 2.0254; 05: 2.2
| Std Dev. |  | 02: 0.11832; 03: 0.13292; 04: 0.035896; 05: 0.18348
| Rel.Std.Dev. |  | 02: 4.733%; 03: 6.935%; 04: 1.771%; 05: 8.036%
| PDM3 |  | 02: 15.42%; 03: -11.51%; 04: -6.413%; 05: 5.421%
| Z-Score (Absolute) |  | 02: 1.27; 03: 0.88; 04: 0.47; 05: 0.47
| Analytical results for Zn in OREAS 24b (Certified Value 105 ppm) | [10 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 0
0
| 1 | 1 | 114 | 108 | 111 | 106 | 96 | 95.5029 | 104 | 105 | 100 | 1
| 1 | 2 | 111 | 106 | 111 | 105 | 99 | 98.2086 | 118 | 103 | 94 | 4
| 1 | 3 | 112 | 106 | 108 | 107 | 101 | 96.1468 | 112 | 107 | 97 | 16
| 1 | 4 | 113 | 104 | 109 | 105 | 104 | 98.0834 | 111 | 103 | 97 | 104.6524
| 1 | 5 | 110 | 108 | 107 | 107 | 97 | 99.8732 | 107 | 103 | 100 | 109
| 1 | 6 | 113 | 108 | 108 | 105 | 97 | 97.4139 | 104 | 107 | 98
| Mean |  | 112.1667 | 106.6667 | 109 | 105.8333 | 99 | 97.5381 | 109.3333 | 104.6667 | 97.6667
| Median |  | 112.5 | 107 | 108.5 | 105.5 | 98 | 97.7487 | 109 | 104 | 97.5
| Std Dev. |  | 1.472 | 1.633 | 1.6733 | 0.98319 | 3.0332 | 1.568 | 5.4283 | 1.9664 | 2.2509
| Rel.Std.Dev. |  | 1.312% | 1.531% | 1.535% | 0.929% | 3.064% | 1.608% | 4.965% | 1.879% | 2.305%
| PDM3 |  | 7.18% | 1.925% | 4.154% | 1.128% | -5.401% | -6.798% | 4.473% | 0.01365% | -6.675%
| Z-Score (Absolute) |  | 1.22 | 0.16 | 0.61 | 0 | 1.32 | 1.6 | 0.67 | 0.22 | 1.57
| Analytical results for Zr in OREAS 24b (Certified Value 134 ppm) | [10 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  | 0
0
| 1 | 1 | 158 | 163.8 | 97.5 | 135.4 | 128.5962 | 110 | 141 | 123.4 |  | 1
| 1 | 2 | 169 | 165.6 | 95.4 | 134.7 | 128.6703 | 119 | 138.5 | 121.2 |  | 0
| 1 | 3 | 172 | 169.2 | 102 | 141.3 | 130.3842 | 111 | 146 | 119.7 |  | 16
| 1 | 4 | 172 | 166.8 | 102 | 149.4 | 127.005 | 93 | 140.5 | 120.2 |  | 134.4032
| 1 | 5 | 169 | 166.3 | 97.5 | 143.8 | 129.4634 | 112 | 135 | 120.9 |  | 110
| 1 | 6 | 169 | 168 | 96.8 | 138.4 | 126.8358 | 151 | 143 | 121.9
| Mean |  | 168.1667 | 166.6167 | 98.5333 | 140.5 | 128.4925 | 116 | 140.6667 | 121.2167
| Median |  | 169 | 166.55 | 97.5 | 139.85 | 128.6332 | 111.5 | 140.75 | 121.05
| Std Dev. |  | 5.1929 | 1.8809 | 2.7926 | 5.5642 | 1.3799 | 19.1833 | 3.7639 | 1.3167
| Rel.Std.Dev. |  | 3.088% | 1.129% | 2.834% | 3.96% | 1.074% | 16.54% | 2.676% | 1.086%
| PDM3 |  | 25.12% | 23.97% | -26.69% | 4.536% | -4.398% | -13.69% | 4.66% | -9.811%
| Z-Score (Absolute) |  | 1.43 | 1.36 | 1.53 | 0.25 | 0.25 | 0.79 | 0.26 | 0.56
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 0.15-50g
| Analytical results for Ag in OREAS 24b (Indicative Value 0.058 ppm) | [10 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 3
3
| 1 | 1 | < 0.2 | < 0.05 | 0.06 | < 2 | 0.08 | 0.0625 | 0.011 | < 0.2 | < 0.1 | 1
| 1 | 2 | < 0.2 | < 0.05 | 0.06 | < 2 | 0.2 | 0.06543 | 0.014 | < 0.2 | < 0.1 | 7
| 1 | 3 | < 0.2 | < 0.05 | 0.06 | < 2 | 0.08 | 0.066667 | 0.027 | < 0.2 | < 0.1 | 16
| 1 | 4 | < 0.2 | < 0.05 | 0.07 | < 2 | 0.08 | 0.064922 | 0.023 | < 0.2 | < 0.1 | 0.057508
| 1 | 5 | < 0.2 | < 0.05 | 0.06 | < 2 | 0.08 | 0.060811 | 0.028 | < 0.2 | < 0.1 | 28
| 1 | 6 | < 0.2 | < 0.05 | 0.06 | < 2 | 0.09 | 0.066872 | 0.028 | < 0.2 | < 0.1
| Mean |  | 03: 0.061667; 05: 0.10167; 06: 0.064534; 07: 0.021833
| Median |  | 03: 0.06; 05: 0.08; 06: 0.065176; 07: 0.025
| Std Dev. |  | 03: 0.0040825; 05: 0.048339; 06: 0.002407; 07: 0.0075211
| Rel.Std.Dev. |  | 03: 6.62%; 05: 47.55%; 06: 3.73%; 07: 34.45%
| PDM3 |  | 03: 7.231%; 05: 76.79%; 06: 12.22%; 07: -62.03%
| Z-Score (Absolute) |  | 0.67 | 0.75 | 0.05 | 17.79 | 0.71 | 0 | 0.81 | 0.67 | 0.28
| Analytical results for Al in OREAS 24b (Certified Value 3.15 wt.%) | [10 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
3
| 1 | 1 | 3.07 | 3.2 | 3.19 | 3.0532 | 3.3367 | 3.9 | 3.3 | 2.92 |  | 1
| 1 | 2 | 3.02 | 3.03 | 3.54 | 3.2487 | 3.3172 | 3.92 | 3.3 | 3 |  | 0
| 1 | 3 | 3.11 | 2.95 | 3.18 | 3.0828 | 3.3908 | 3.96 | 3.4 | 3.01 |  | 16
| 1 | 4 | 3.19 | 3.03 | 3.21 | 3.1737 | 3.3613 | 3.82 | 3.38 | 2.96 |  | 3.152
| 1 | 5 | 3.05 | 2.92 | 3.15 | 2.9654 | 3.34 | 4.04 | 3.33 | 3.01 |  | 112
| 1 | 6 | 3.04 | 2.99 | 3.24 | 3.146 | 3.3351 | 3.99 | 3.3 | 2.96
| Mean |  | 3.08 | 3.02 | 3.2517 | 3.1116 | 3.3468 | 3.9383 | 3.335 | 2.9767
| Median |  | 3.06 | 3.01 | 3.2 | 3.1144 | 3.3383 | 3.94 | 3.315 | 2.98
| Std Dev. |  | 0.061968 | 0.098387 | 0.14442 | 0.099449 | 0.025726 | 0.076529 | 0.044609 | 0.036148
| Rel.Std.Dev. |  | 2.012% | 3.258% | 4.441% | 3.196% | 0.7687% | 1.943% | 1.338% | 1.214%
| PDM3 |  | -2.285% | -4.188% | 3.161% | -1.281% | 6.181% | 24.95% | 5.805% | -5.563%
| Z-Score (Absolute) |  | 0.44 | 0.69 | 0.3 | 0.3 | 0.71 | 3.24 | 0.66 | 0.88
| Analytical results for As in OREAS 24b (Certified Value 7.96 ppm) | [10 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 2
3
| 1 | 1 | 6 | 8 | 7.9 | 9 | 7.6 | 8.7015 | 9.7 | 8 | 7 | 1
| 1 | 2 | 8 | 8 | 7.8 | 10 | 8.3 | 8.1817 | 9.2 | 9 | 8.7 | 3
| 1 | 3 | 8 | 8.2 | 8.2 | 11 | 7.7 | 8.5376 | 9.4 | 8 | 7.7 | 16
| 1 | 4 | 7 | 8.2 | 8.1 | 10 | 8.1 | 7.9081 | 8.9 | 7 | 7.7 | 7.9567
| 1 | 5 | 6 | 8.2 | 7.9 | 11 | 7.7 | 9.2745 | 9.8 | 9 | 6.8 | 113
| 1 | 6 | 10 | 8.2 | 8 | 11 | 7.9 | 8.1787 | 9.4 | 8 | 7.5
| Mean |  | 7.5 | 8.1333 | 7.9833 | 10.3333 | 7.8833 | 8.4637 | 9.4 | 8.1667 | 7.5667
| Median |  | 7.5 | 8.2 | 7.95 | 10.5 | 7.8 | 8.3597 | 9.4 | 8 | 7.6
| Std Dev. |  | 1.5166 | 0.10328 | 0.1472 | 0.8165 | 0.27142 | 0.48761 | 0.32863 | 0.75277 | 0.66833
| Rel.Std.Dev. |  | 20.22% | 1.27% | 1.844% | 7.902% | 3.443% | 5.761% | 3.496% | 9.218% | 8.833%
| PDM3 |  | -5.74% | 2.22% | 0.3345% | 29.87% | -0.9223% | 6.372% | 18.14% | 2.639% | -4.902%
| Z-Score (Absolute) |  | 1.29 | 0 | 0.31 | 4.49 | 0.51 | 0.67 | 2.59 | 0.07 | 1.16
| Analytical results for Au in OREAS 24b (Indicative Value 0.002 ppm) | [10 empty] | Indicative
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 07 | 09 | 10 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*GFAAS | AR*MS | AR*MS | AR*MS |  |  | 3
3
| 1 | 1 | 0.003 | 0.002 | 0.00192 | 0.002 | < 0.005 | 0.003 | 0.003 |  |  | 1
| 1 | 2 | 0.005 | 0.001 | 0.00158 | 0.003 | < 0.005 | 0.0015 | 0.002 |  |  | 23
| 1 | 3 | 0.004 | 0.001 | 0.00219 | 0.002 | < 0.005 | 0.0014 | 0.003 |  |  | 16
| 1 | 4 | 0.004 | 0.002 | 0.00265 | 0.002 | < 0.005 | < 0.0005 | 0.003 |  |  | 0.0023407
| 1 | 5 | 0.003 | 0.001 | 0.00192 | 0.003 | 0.091 | 0.0019 | 0.003 |  |  | 29
| 1 | 6 | 0.003 | 0.001 | 0.00275 | 0.002 | < 0.005 | < 0.0005 | 0.003
| Mean |  | 0.0036667 | 0.0013333 | 0.0021683 | 0.0023333 | 0.091 | 0.00195 | 0.0028333
| Median |  | 0.0035 | 0.001 | 0.002055 | 0.002 | 0.091 | 0.0017 | 0.003
| Std Dev. |  | 0.0008165 | 0.0005164 | 0.00045613 | 0.0005164 |  | 0.00073258 | 0.00040825
| Rel.Std.Dev. |  | 22.27% | 38.73% | 21.04% | 22.13% |  | 37.57% | 14.41%
| PDM3 |  | 56.65% | -43.04% | -7.364% | -0.3153% | 3788% | -16.69% | 21.05%
| Z-Score (Absolute) |  | 0.95 | 0.71 | 0.12 | 0 | 10.59 | 0.67 | 0.35
| Analytical results for B in OREAS 24b (Indicative Value 6.23 ppm) | [10 empty] | Indicative
B | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 01; 02: 06; 03: 07; 04: 08; 05: 09 | ppm
|  |  | 01: AR*OES; 02: AR*OES/MS; 03: AR*MS; 04: AR*OES; 05: AR*MS | 2
2
| 1 | 1 | 01: < 10; 02: 6.4808; 03: 5; 04: < 10; 05: < 20 | 1
| 1 | 2 | 01: < 10; 02: 7.3682; 03: 4; 04: < 10; 05: < 20 | 9
| 1 | 3 | 01: < 10; 02: 7.3667; 03: 4; 04: < 10; 05: < 20 | 16
| 1 | 4 | 01: < 10; 02: 8.7907; 03: 4; 04: < 10; 05: < 20 | 6.2269
| 1 | 5 | 01: < 10; 02: 6.9923; 03: 4; 04: < 10; 05: < 20 | 30
| 1 | 6 | 01: < 10; 02: 12.7243; 03: 4; 04: < 10; 05: < 20
| Mean |  | 02: 8.2872; 03: 4.1667
| Median |  | 02: 7.3674; 03: 4
| Std Dev. |  | 02: 2.3053; 03: 0.40825
| Rel.Std.Dev. |  | 02: 27.82%; 03: 9.798%
| PDM3 |  | 02: 33.09%; 03: -33.09%
| Z-Score (Absolute) |  | 01: 0; 02: 2.66; 03: 0.67; 04: 0; 05: 4.05
| Analytical results for Ba in OREAS 24b (Certified Value 146 ppm) | [10 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 0
0
| 1 | 1 | 140 | 192.8 | 146 | 139 | 154.9899 | 178 | 150 | 142 |  | 1
| 1 | 2 | 140 | 190.6 | 149 | 146 | 153.4375 | 180 | 150 | 146 |  | 0
| 1 | 3 | 140 | 188.4 | 145 | 137 | 157.4412 | 179 | 150 | 146 |  | 16
| 1 | 4 | 140 | 194.7 | 143 | 149 | 153.3721 | 183 | 150 | 139 |  | 145.7952
| 1 | 5 | 140 | 187.9 | 141 | 133 | 155.7336 | 172 | 150 | 146 |  | 114
| 1 | 6 | 140 | 189.5 | 146 | 141 | 157.6543 | 182 | 150 | 142
| Mean |  | 140 | 190.65 | 145 | 140.8333 | 155.4381 | 179 | 150 | 143.5
| Median |  | 140 | 190.05 | 145.5 | 140 | 155.3618 | 179.5 | 150 | 144
| Std Dev. |  | 0 | 2.6448 | 2.7568 | 5.8793 | 1.8699 | 3.8987 | 0 | 2.9496
| Rel.Std.Dev. |  | 0% | 1.387% | 1.901% | 4.175% | 1.203% | 2.178% | 0% | 2.055%
| PDM3 |  | -3.975% | 30.77% | -0.5454% | -3.403% | 6.614% | 22.77% | 2.884% | -1.574%
| Z-Score (Absolute) |  | 0.71 | 4.11 | 0.24 | 0.63 | 0.76 | 3 | 0.24 | 0.38
| Analytical results for Be in OREAS 24b (Indicative Value 1.65 ppm) | [10 empty] | Indicative
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 08 |  |  |  | ppm
|  |  | AR*OES | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  |  |  | 2
2
| 1 | 1 | 1.3 | 1.2 | 1.7 | 1.98 | 2.4 | 1.4 |  |  |  | 1
| 1 | 2 | 1.3 | 1.1 | 1.7 | 1.99 | 2.5 | 1.4 |  |  |  | 25
| 1 | 3 | 1.4 | 1 | 1.8 | 2.1 | 2.1 | 1.4 |  |  |  | 16
| 1 | 4 | 1.4 | 1.2 | 1.8 | 2.02 | 2.1 | 1.4 |  |  |  | 1.6547
| 1 | 5 | 1.4 | 1.2 | 1.7 | 2 | 2 | 1.4 |  |  |  | 31
| 1 | 6 | 1.4 | 1.2 | 1.8 | 2 | 2.2 | 1.4
| Mean |  | 1.3667 | 1.15 | 1.75 | 2.015 | 2.2167 | 1.4
| Median |  | 1.4 | 1.2 | 1.75 | 2 | 2.15 | 1.4
| Std Dev. |  | 0.05164 | 0.083666 | 0.054772 | 0.043704 | 0.19408 | 2.4324e-16
| Rel.Std.Dev. |  | 3.779% | 7.275% | 3.13% | 2.169% | 8.755% | 1.737e-14%
| PDM3 |  | -17.41% | -30.5% | 5.758% | 21.77% | 33.96% | -15.39%
| Z-Score (Absolute) |  | 0.44 | 0.9 | 0.37 | 0.94 | 1.37 | 0.37
| Analytical results for Bi in OREAS 24b (Certified Value 0.73 ppm) | [10 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
3
| 1 | 1 | 2 | 0.76 | 0.73 | 0.71 | 0.70665 | 0.74 | < 2 | 0.6 |  | 1
| 1 | 2 | 2 | 0.74 | 0.76 | 0.94 | 0.68555 | 0.74 | < 2 | 0.6 |  | 10
| 1 | 3 | 2 | 0.7 | 0.75 | 0.8 | 0.69314 | 0.87 | 2 | 0.7 |  | 16
| 1 | 4 | 2 | 0.7 | 0.74 | 0.71 | 0.70349 | 0.77 | < 2 | 0.6 |  | 0.73357
| 1 | 5 | 3 | 0.74 | 0.73 | 0.75 | 0.80405 | 0.78 | < 2 | 0.7 |  | 115
| 1 | 6 | 2 | 0.72 | 0.75 | 0.74 | 0.70576 | 0.9 | 3 | 1.1
| Mean |  | 2.1667 | 0.72667 | 0.74333 | 0.775 | 0.71644 | 0.8 | 2.5 | 0.71667
| Median |  | 2 | 0.73 | 0.745 | 0.745 | 0.70462 | 0.775 | 2.5 | 0.65
| Std Dev. |  | 0.40825 | 0.024221 | 0.012111 | 0.08735 | 0.043706 | 0.068411 | 0.70711 | 0.19408
| Rel.Std.Dev. |  | 18.84% | 3.333% | 1.629% | 11.27% | 6.1% | 8.551% | 28.28% | 27.08%
| PDM3 |  | 195.4% | -0.9415% | 1.33% | 5.647% | -2.336% | 9.055% | 240.8% | -2.305%
| Z-Score (Absolute) |  | 22.78 | 0.53 | 0.26 | 0.26 | 0.69 | 0.66 | 11.99 | 0.69
| Analytical results for Ca in OREAS 24b (Certified Value 0.461 wt.%) | [10 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 3
3
| 1 | 1 | 0.48 | 0.49 | 0.46 | 0.45 | 0.46081 | 0.55 | 0.48 | 0.42 |  | 1
| 1 | 2 | 0.47 | 0.48 | 0.49 | 0.47 | 0.46196 | 0.58 | 0.47 | 0.46 |  | 0
| 1 | 3 | 0.48 | 0.46 | 0.44 | 0.45 | 0.46457 | 0.54 | 0.48 | 0.44 |  | 16
| 1 | 4 | 0.49 | 0.48 | 0.45 | 0.47 | 0.46424 | 0.55 | 0.47 | 0.41 |  | 0.46078
| 1 | 5 | 0.47 | 0.47 | 0.44 | 0.44 | 0.46023 | 0.52 | 0.47 | 0.44 |  | 116
| 1 | 6 | 0.47 | 0.47 | 0.44 | 0.45 | 0.46507 | 0.6 | 0.47 | 0.45
| Mean |  | 0.47667 | 0.475 | 0.45333 | 0.455 | 0.46281 | 0.55667 | 0.47333 | 0.43667
| Median |  | 0.475 | 0.475 | 0.445 | 0.45 | 0.4631 | 0.55 | 0.47 | 0.44
| Std Dev. |  | 0.008165 | 0.010488 | 0.019664 | 0.012247 | 0.0020807 | 0.028752 | 0.005164 | 0.018619
| Rel.Std.Dev. |  | 1.713% | 2.208% | 4.338% | 2.692% | 0.4496% | 5.165% | 1.091% | 4.264%
| PDM3 |  | 3.447% | 3.085% | -1.617% | -1.255% | 0.4408% | 20.81% | 2.724% | -5.234%
| Z-Score (Absolute) |  | 0.53 | 0.43 | 0.92 | 0.81 | 0.33 | 5.51 | 0.33 | 1.95
| Analytical results for Cd in OREAS 24b (Indicative Value 0.046 ppm) | [10 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 3
3
| 1 | 1 | < 0.5 | < 0.05 | 0.07 | < 0.5 | < 0.05 | 0.050403 | 0.05 | < 0.5 | < 0.1 | 1
| 1 | 2 | < 0.5 | < 0.05 | 0.06 | < 0.5 | 0.08 | 0.043945 | 0.05 | < 0.5 | < 0.1 | 11
| 1 | 3 | < 0.5 | 0.05 | 0.06 | < 0.5 | < 0.05 | 0.034314 | 0.04 | < 0.5 | < 0.1 | 16
| 1 | 4 | < 0.5 | < 0.05 | 0.05 | < 0.5 | 0.06 | 0.037791 | 0.03 | < 0.5 | < 0.1 | 0.045875
| 1 | 5 | < 0.5 | 0.05 | 0.06 | < 0.5 | 0.06 | 0.043436 | 0.05 | < 0.5 | < 0.1 | 32
| 1 | 6 | < 0.5 | < 0.05 | 0.06 | < 0.5 | 0.06 | 0.048354 | 0.04 | < 0.5 | < 0.1
| Mean |  | 02: 0.05; 03: 0.06; 05: 0.065; 06: 0.043041; 07: 0.043333
| Median |  | 02: 0.05; 03: 0.06; 05: 0.06; 06: 0.043691; 07: 0.045
| Std Dev. |  | 02: 0; 03: 0.0063246; 05: 0.01; 06: 0.0061178; 07: 0.008165
| Rel.Std.Dev. |  | 02: 0%; 03: 10.54%; 05: 15.38%; 06: 14.21%; 07: 18.84%
| PDM3 |  | 02: 8.992%; 03: 30.79%; 05: 41.69%; 06: -6.178%; 07: -5.54%
| Z-Score (Absolute) |  | 15.5 | 1.43 | 0.65 | 15.5 | 0 | 0.67 | 0.65 | 15.5 | 0.13
| Analytical results for Ce in OREAS 24b (Indicative Value 61 ppm) | [10 empty] | Indicative
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 |  |  |  | ppm
|  |  | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 0
0
| 1 | 1 | 41.2 | 78.89 | 53.9 | 50.63 | 73.9395 | 68.4 |  |  |  | 1
| 1 | 2 | 40.1 | 78.29 | 54.6 | 53.24 | 73.7334 | 69 |  |  |  | 27
| 1 | 3 | 39.5 | 76.71 | 52.3 | 53.15 | 74.1765 | 68 |  |  |  | 16
| 1 | 4 | 37.9 | 79.92 | 59.9 | 56.92 | 73.4535 | 66.4 |  |  |  | 61.2619
| 1 | 5 | 39.9 | 76.82 | 54.8 | 52.16 | 72.2297 | 68 |  |  |  | 33
| 1 | 6 | 42 | 79.03 | 54.2 | 54.68 | 72.8961 | 70.4
| Mean |  | 40.1 | 78.2767 | 54.95 | 53.4633 | 73.4048 | 68.3667
| Median |  | 40 | 78.59 | 54.4 | 53.195 | 73.5934 | 68.2
| Std Dev. |  | 1.4184 | 1.2822 | 2.5821 | 2.1589 | 0.72544 | 1.3171
| Rel.Std.Dev. |  | 3.537% | 1.638% | 4.699% | 4.038% | 0.9883% | 1.926%
| PDM3 |  | -34.54% | 27.77% | -10.3% | -12.73% | 19.82% | 11.6%
| Z-Score (Absolute) |  | 1.46 | 1.12 | 0.45 | 0.55 | 0.79 | 0.45
| Analytical results for Co in OREAS 24b (Certified Value 15.7 ppm) | [10 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*OES | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
2
| 1 | 1 | 14 | 14 | 16.8 | 15.3 | 15 | 16.86 | 17.1 | 15 | 14.6 | 1
| 1 | 2 | 15 | 14 | 16.3 | 15.7 | 16.2 | 16.97 | 18.1 | 14 | 15.7 | 5
| 1 | 3 | 15 | 14 | 16.4 | 15.4 | 15.1 | 17.44 | 17.9 | 15 | 15.2 | 16
| 1 | 4 | 15 | 14 | 17 | 16.3 | 16.3 | 18.33 | 17.3 | 14 | 14.8 | 15.6533
| 1 | 5 | 15 | 14 | 16.4 | 15.1 | 14.5 | 17.44 | 16.7 | 14 | 14.5 | 117
| 1 | 6 | 15 | 15 | 16.7 | 15.7 | 15.6 | 17.44 | 18.6 | 14 | 14.5
| Mean |  | 14.8333 | 14.1667 | 16.6 | 15.5833 | 15.45 | 17.4133 | 17.6167 | 14.3333 | 14.8833
| Median |  | 15 | 14 | 16.55 | 15.55 | 15.35 | 17.44 | 17.6 | 14 | 14.7
| Std Dev. |  | 0.40825 | 0.40825 | 0.27568 | 0.42151 | 0.71204 | 0.51868 | 0.70545 | 0.5164 | 0.47924
| Rel.Std.Dev. |  | 2.752% | 2.882% | 1.661% | 2.705% | 4.609% | 2.979% | 4.004% | 3.603% | 3.22%
| PDM3 |  | -5.239% | -9.497% | 6.048% | -0.4472% | -1.299% | 11.24% | 12.54% | -8.433% | -4.919%
| Z-Score (Absolute) |  | 0.37 | 0.77 | 0.69 | 0.08 | 0 | 1.19 | 1.31 | 0.67 | 0.34
| Analytical results for Cr in OREAS 24b (Certified Value 106 ppm) | [10 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 0
0
| 1 | 1 | 103 | 110 | 94 | 102 | 103.9113 | 120 | 106 | 106 |  | 1
| 1 | 2 | 103 | 105 | 96 | 112 | 109.6094 | 119 | 107 | 116 |  | 0
| 1 | 3 | 105 | 105 | 93 | 104 | 105.8529 | 116 | 107 | 111 |  | 16
| 1 | 4 | 111 | 105 | 95 | 108 | 107.9748 | 112 | 106 | 105 |  | 106.4188
| 1 | 5 | 104 | 105 | 92 | 101 | 111.4189 | 117 | 108 | 108 |  | 118
| 1 | 6 | 106 | 105 | 93 | 106 | 111.7078 | 112 | 109 | 107
| Mean |  | 105.3333 | 105.8333 | 93.8333 | 105.5 | 108.4125 | 116 | 107.1667 | 108.8333
| Median |  | 104.5 | 105 | 93.5 | 105 | 108.7921 | 116.5 | 107 | 107.5
| Std Dev. |  | 3.0111 | 2.0412 | 1.472 | 4.0866 | 3.1091 | 3.4059 | 1.169 | 4.0702
| Rel.Std.Dev. |  | 2.859% | 1.929% | 1.569% | 3.874% | 2.868% | 2.936% | 1.091% | 3.74%
| PDM3 |  | -1.02% | -0.5501% | -11.83% | -0.8633% | 1.874% | 9.003% | 0.7028% | 2.269%
| Z-Score (Absolute) |  | 0.51 | 0.29 | 5.55 | 0.44 | 0.84 | 4.16 | 0.29 | 1.02
| Analytical results for Cs in OREAS 24b (Certified Value 9.15 ppm) | [10 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 |  |  |  | ppm
|  |  | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
3
| 1 | 1 | 8.94 | 12 | 9.4 | 8.7 | 9.7812 | 8.37 |  |  |  | 1
| 1 | 2 | 8.56 | 12 | 10.1 | 9.38 | 9.6338 | 8.61 |  |  |  | 0
| 1 | 3 | 8.58 | 11.8 | 9.5 | 8.82 | 9.8647 | 8.54 |  |  |  | 16
| 1 | 4 | 8.64 | 12.2 | 10.3 | 9.41 | 9.8343 | 8.65 |  |  |  | 9.1493
| 1 | 5 | 8.6 | 11.7 | 9.7 | 8.53 | 9.6863 | 8.34 |  |  |  | 119
| 1 | 6 | 8.68 | 11.9 | 9.9 | 9.21 | 9.8879 | 8.66
| Mean |  | 8.6667 | 11.9333 | 9.8167 | 9.0083 | 9.7814 | 8.5283
| Median |  | 8.62 | 11.95 | 9.8 | 9.015 | 9.8078 | 8.575
| Std Dev. |  | 0.14067 | 0.17512 | 0.34881 | 0.37403 | 0.10189 | 0.14106
| Rel.Std.Dev. |  | 1.623% | 1.467% | 3.553% | 4.152% | 1.042% | 1.654%
| PDM3 |  | -5.275% | 30.43% | 7.294% | -1.541% | 6.908% | -6.787%
| Z-Score (Absolute) |  | 0.85 | 2.98 | 0.49 | 0.45 | 0.45 | 1.02
| Analytical results for Cu in OREAS 24b (Certified Value 36.4 ppm) | [10 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*OES | AR*MS | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
2
| 1 | 1 | 36 | 41 | 37.5 | 35 | 35.5 | 35.0887 | 41.5 | 37 | 35.6 | 1
| 1 | 2 | 36 | 38 | 36.7 | 40.5 | 38.2 | 35.6719 | 39.5 | 37 | 39.3 | 12
| 1 | 3 | 36 | 38 | 37 | 34.3 | 35.2 | 36.1127 | 40.2 | 38 | 37.7 | 16
| 1 | 4 | 38 | 39 | 38.1 | 33.9 | 36.5 | 35.2112 | 41.2 | 37 | 34.9 | 36.3669
| 1 | 5 | 36 | 37.5 | 37 | 33.8 | 34.3 | 36.6303 | 42.2 | 38 | 34.8 | 120
| 1 | 6 | 37 | 38 | 37.4 | 34.8 | 35.8 | 36.7377 | 46.7 | 36 | 35.5
| Mean |  | 36.5 | 38.5833 | 37.2833 | 35.3833 | 35.9167 | 35.9088 | 41.8833 | 37.1667 | 36.3
| Median |  | 36 | 38 | 37.2 | 34.55 | 35.65 | 35.8923 | 41.35 | 37 | 35.55
| Std Dev. |  | 0.83666 | 1.2813 | 0.49565 | 2.5514 | 1.3318 | 0.70206 | 2.5467 | 0.75277 | 1.8055
| Rel.Std.Dev. |  | 2.292% | 3.321% | 1.329% | 7.211% | 3.708% | 1.955% | 6.08% | 2.025% | 4.974%
| PDM3 |  | 0.3659% | 6.095% | 2.52% | -2.705% | -1.238% | -1.26% | 15.17% | 2.199% | -0.184%
| Z-Score (Absolute) |  | 0 | 2.11 | 0.79 | 1.13 | 0.59 | 0.6 | 5.45 | 0.67 | 0.2
| Analytical results for Dy in OREAS 24b (Indicative Value 2.65 ppm) | [10 empty] | Indicative
Dy | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 2
2
| 1 | 1 | 01: 2.29; 02: 2.98 | 1
| 1 | 2 | 01: 2.25; 02: 3.07 | 17
| 1 | 3 | 01: 2.2; 02: 3.1 | 16
| 1 | 4 | 01: 2.19; 02: 3.02 | 2.6475
| 1 | 5 | 01: 2.27; 02: 3.05 | 34
| 1 | 6 | 01: 2.3; 02: 3.05
| Mean |  | 01: 2.25; 02: 3.045
| Median |  | 01: 2.26; 02: 3.05
| Std Dev. |  | 01: 0.046043; 02: 0.041352
| Rel.Std.Dev. |  | 01: 2.046%; 02: 1.358%
| PDM3 |  | 01: -15.01%; 02: 15.01%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Er in OREAS 24b (Indicative Value 1.21 ppm) | [10 empty] | Indicative
Er | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 2
2
| 1 | 1 | 01: 1.14; 02: 1.3 | 1
| 1 | 2 | 01: 1.12; 02: 1.3 | 18
| 1 | 3 | 01: 1.09; 02: 1.3 | 16
| 1 | 4 | 01: 1.11; 02: 1.3 | 1.2083
| 1 | 5 | 01: 1.12; 02: 1.3 | 35
| 1 | 6 | 01: 1.12; 02: 1.3
| Mean |  | 01: 1.1167; 02: 1.3
| Median |  | 01: 1.12; 02: 1.3
| Std Dev. |  | 01: 0.01633; 02: 0
| Rel.Std.Dev. |  | 01: 1.462%; 02: 0%
| PDM3 |  | 01: -7.586%; 02: 7.586%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Eu in OREAS 24b (Indicative Value 0.66 ppm) | [10 empty] | Indicative
Eu | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 2
2
| 1 | 1 | 01: 0.532; 02: 0.8 | 1
| 1 | 2 | 01: 0.512; 02: 0.8 | 19
| 1 | 3 | 01: 0.522; 02: 0.8 | 16
| 1 | 4 | 01: 0.494; 02: 0.8 | 0.65833
| 1 | 5 | 01: 0.508; 02: 0.8 | 36
| 1 | 6 | 01: 0.532; 02: 0.8
| Mean |  | 01: 0.51667; 02: 0.8
| Median |  | 01: 0.517; 02: 0.8
| Std Dev. |  | 01: 0.014895; 02: 1.2162e-16
| Rel.Std.Dev. |  | 01: 2.883%; 02: 1.52e-14%
| PDM3 |  | 01: -21.52%; 02: 21.52%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Fe in OREAS 24b (Certified Value 3.93 wt.%) | [10 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
3
| 1 | 1 | 3.87 | 4.39 | 3.71 | 3.72 | 3.9554 | 4.21 | 3.88 | 3.55 |  | 1
| 1 | 2 | 3.8 | 4.24 | 3.74 | 4.02 | 3.9439 | 4.31 | 3.85 | 3.68 |  | 0
| 1 | 3 | 3.83 | 4.14 | 3.65 | 3.79 | 4.0026 | 4.48 | 3.92 | 3.71 |  | 16
| 1 | 4 | 3.96 | 4.29 | 3.7 | 3.95 | 3.9611 | 4.4 | 3.86 | 3.61 |  | 3.9257
| 1 | 5 | 3.8 | 4.12 | 3.61 | 3.7 | 3.9556 | 4.33 | 3.85 | 3.6 |  | 121
| 1 | 6 | 3.79 | 4.2 | 3.7 | 3.91 | 3.9957 | 4.47 | 3.81 | 3.61
| Mean |  | 3.8417 | 4.23 | 3.685 | 3.8483 | 3.9691 | 4.3667 | 3.8617 | 3.6267
| Median |  | 3.815 | 4.22 | 3.7 | 3.85 | 3.9584 | 4.365 | 3.855 | 3.61
| Std Dev. |  | 0.064936 | 0.1004 | 0.046797 | 0.13075 | 0.024069 | 0.10367 | 0.03656 | 0.058195
| Rel.Std.Dev. |  | 1.69% | 2.374% | 1.27% | 3.398% | 0.6064% | 2.374% | 0.9468% | 1.605%
| PDM3 |  | -2.14% | 7.752% | -6.131% | -1.97% | 1.106% | 11.23% | -1.631% | -7.617%
| Z-Score (Absolute) |  | 0.06 | 1.78 | 0.81 | 0.03 | 0.54 | 2.43 | 0.03 | 1.08
| Analytical results for Ga in OREAS 24b (Certified Value 10.8 ppm) | [10 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
1
| 1 | 1 | 10 | 11 | 12.5 | 10 | 10.5 | 12.71 | 14.1 | 10 | 10 | 1
| 1 | 2 | 10 | 10.2 | 12.5 | 10 | 10.9 | 12.71 | 14.6 | 10 | 10 | 0
| 1 | 3 | 10 | 10.8 | 12.4 | 10 | 10.5 | 12.25 | 14.4 | 10 | 9 | 16
| 1 | 4 | 10 | 10.4 | 12.8 | 11 | 11.3 | 12.45 | 12.3 | 10 | 9 | 10.765
| 1 | 5 | 10 | 10.6 | 12.4 | 10 | 10.3 | 12.21 | 13.3 | 10 | 9 | 122
| 1 | 6 | 10 | 10.6 | 12.6 | 11 | 10.8 | 12.29 | 13.5 | 10 | 10
| Mean |  | 10 | 10.6 | 12.5333 | 10.3333 | 10.7167 | 12.4367 | 13.7 | 10 | 9.5
| Median |  | 10 | 10.6 | 12.5 | 10 | 10.65 | 12.37 | 13.8 | 10 | 9.5
| Std Dev. |  | 0 | 0.28284 | 0.15055 | 0.5164 | 0.36009 | 0.22686 | 0.85088 | 0 | 0.54772
| Rel.Std.Dev. |  | 0% | 2.668% | 1.201% | 4.997% | 3.36% | 1.824% | 6.211% | 0% | 5.766%
| PDM3 |  | -7.106% | -1.533% | 16.43% | -4.01% | -0.449% | 15.53% | 27.26% | -7.106% | -11.75%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.17 | 0.3 | 0.13 | 2.06 | 3.48 | 0.67 | 1.24
| Analytical results for Gd in OREAS 24b (Indicative Value 3.96 ppm) | [10 empty] | Indicative
Gd | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 2
2
| 1 | 1 | 01: 3.27; 02: 4.6 | 1
| 1 | 2 | 01: 3.26; 02: 4.5 | 20
| 1 | 3 | 01: 3.16; 02: 4.9 | 16
| 1 | 4 | 01: 3.09; 02: 4.7 | 3.9633
| 1 | 5 | 01: 3.18; 02: 4.9 | 37
| 1 | 6 | 01: 3.3; 02: 4.7
| Mean |  | 01: 3.21; 02: 4.7167
| Median |  | 01: 3.22; 02: 4.7
| Std Dev. |  | 01: 0.08; 02: 0.16021
| Rel.Std.Dev. |  | 01: 2.492%; 02: 3.397%
| PDM3 |  | 01: -19.01%; 02: 19.01%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Ge in OREAS 24b (Indicative Value 0.26 ppm) | [10 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 04; 02: 06; 03: 07 | ppm
|  |  | 01: 2A*OES/MS; 02: AR*OES/MS; 03: AR*MS | 2
2
| 1 | 1 | 01: < 0.1; 02: 0.26512; 03: 0.3 | 1
| 1 | 2 | 01: < 0.1; 02: 0.20898; 03: 0.3 | 32
| 1 | 3 | 01: < 0.1; 02: 0.19902; 03: 0.3 | 16
| 1 | 4 | 01: < 0.1; 02: 0.19767; 03: 0.3 | 0.25665
| 1 | 5 | 01: < 0.1; 02: 0.19402; 03: 0.3 | 38
| 1 | 6 | 01: < 0.1; 02: 0.21502; 03: 0.3
| Mean |  | 02: 0.21331; 03: 0.3
| Median |  | 02: 0.204; 03: 0.3
| Std Dev. |  | 02: 0.026559; 03: 0
| Rel.Std.Dev. |  | 02: 12.45%; 03: 0%
| PDM3 |  | 02: -16.89%; 03: 16.89%
| Z-Score (Absolute) |  | 01: 1.27; 02: 0; 03: 0.67
| Analytical results for Hf in OREAS 24b (Indicative Value 0.52 ppm) | [10 empty] | Indicative
Hf | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 07 | ppm
|  |  | 01: AR*MS; 02: 2A*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS | 2
2
| 1 | 1 | 01: 0.6; 02: 0.6; 03: 0.54; 04: 0.97379; 05: 0.3 | 1
| 1 | 2 | 01: 0.6; 02: < 0.5; 03: 0.65; 04: 0.95996; 05: 0.4 | 33
| 1 | 3 | 01: 0.6; 02: < 0.5; 03: 0.59; 04: 0.95392; 05: 0.4 | 16
| 1 | 4 | 01: 0.6; 02: 0.6; 03: 0.51; 04: 0.95252; 05: 0.4 | 0.52292
| 1 | 5 | 01: 0.6; 02: 0.6; 03: 0.58; 04: 0.95077; 05: 0.6 | 39
| 1 | 6 | 01: 0.6; 02: 0.6; 03: 0.58; 04: 0.93313; 05: 0.5
| Mean |  | 01: 0.6; 02: 0.6; 03: 0.575; 04: 0.95402; 05: 0.43333
| Median |  | 01: 0.6; 02: 0.6; 03: 0.58; 04: 0.95322; 05: 0.4
| Std Dev. |  | 01: 0; 02: 0; 03: 0.047645; 04: 0.013232; 05: 0.10328
| Rel.Std.Dev. |  | 01: 0%; 02: 0%; 03: 8.286%; 04: 1.387%; 05: 23.83%
| PDM3 |  | 01: 14.74%; 02: 14.74%; 03: 9.96%; 04: 82.44%; 05: -17.13%
| Z-Score (Absolute) |  | 01: 0.18; 02: 0.67; 03: 0; 04: 2.79; 05: 1.04
| Analytical results for Hg in OREAS 24b (Indicative Value < 0.01 ppm) | [10 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 06 | 08 | 09 |  |  |  | ppm
|  |  | AR*OES | AR*MS | 2A*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 3
3
| 1 | 1 | < 1 | < 0.01 | < 3 | 0.013105 | < 1 | < 0.01 |  |  |  | 1
| 1 | 2 | < 1 | < 0.01 | < 3 | 0.014648 | < 1 | < 0.01 |  |  |  | 22
| 1 | 3 | < 1 | < 0.01 | < 3 | 0.011765 | < 1 | < 0.01 |  |  |  | 16
| 1 | 4 | < 1 | < 0.01 | < 3 | 0.039729 | 1 | < 0.01 |  |  |  | < 0.01
| 1 | 5 | < 1 | < 0.01 | < 3 | 0.01834 | < 1 | < 0.01 |  |  |  | 28
| 1 | 6 | < 1 | 0.01 | < 3 | 0.019547 | < 1 | < 0.01
| Mean |  | 02: 0.01; 04: 0.019522; 05: 1
| Median |  | 02: 0.01; 04: 0.016494; 05: 1
| Std Dev. |  | 04: 0.010342
| Rel.Std.Dev. |  | 04: 52.98%
PDM3
| Z-Score (Absolute) |  | 0.64 | 0.67 | 3.29 | 0.64 | 0.86 | 0.68
| Analytical results for Ho in OREAS 24b (Indicative Value 0.46 ppm) | [10 empty] | Indicative
Ho | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 2
2
| 1 | 1 | 01: 0.43; 02: 0.5 | 1
| 1 | 2 | 01: 0.42; 02: 0.5 | 15
| 1 | 3 | 01: 0.41; 02: 0.5 | 16
| 1 | 4 | 01: 0.405; 02: 0.5 | 0.45833
| 1 | 5 | 01: 0.41; 02: 0.5 | 29
| 1 | 6 | 01: 0.425; 02: 0.5
| Mean |  | 01: 0.41667; 02: 0.5
| Median |  | 01: 0.415; 02: 0.5
| Std Dev. |  | 01: 0.0098319; 02: 0
| Rel.Std.Dev. |  | 01: 2.36%; 02: 0%
| PDM3 |  | 01: -9.091%; 02: 9.091%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for In in OREAS 24b (Indicative Value 0.048 ppm) | [10 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 |  |  |  | ppm
|  |  | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 3
3
| 1 | 1 | 0.05 | < 0.5 | < 0.1 | < 0.05 | 0.048387 | 0.05 |  |  |  | 1
| 1 | 2 | 0.045 | < 0.5 | < 0.1 | < 0.05 | 0.048828 | 0.05 |  |  |  | 8
| 1 | 3 | 0.045 | < 0.5 | < 0.1 | < 0.05 | 0.04902 | 0.04 |  |  |  | 16
| 1 | 4 | 0.05 | < 0.5 | < 0.1 | < 0.05 | 0.050388 | 0.05 |  |  |  | 0.047627
| 1 | 5 | 0.045 | < 0.5 | < 0.1 | < 0.05 | 0.049228 | 0.04 |  |  |  | 30
| 1 | 6 | 0.045 | < 0.5 | < 0.1 | < 0.05 | 0.05144 | 0.05
| Mean |  | 02: 0.046667; 06: 0.049548; 07: 0.046667
| Median |  | 02: 0.045; 06: 0.049124; 07: 0.05
| Std Dev. |  | 02: 0.002582; 06: 0.0011431; 07: 0.005164
| Rel.Std.Dev. |  | 02: 5.533%; 06: 2.307%; 07: 11.07%
| PDM3 |  | 02: -2.017%; 06: 4.034%; 07: -2.017%
| Z-Score (Absolute) |  | 0.58 | 81.68 | 0.77 | 9.35 | 0.58 | 0.58
| Analytical results for K in OREAS 24b (Certified Value 1.17 wt.%) | [10 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
3
| 1 | 1 | 1.06 | 1.27 | 1.26 | 1.1324 | 1.173 | 1.19 | 1.16 | 1.12 |  | 1
| 1 | 2 | 1.05 | 1.21 | 1.27 | 1.2242 | 1.167 | 1.25 | 1.16 | 1.16 |  | 0
| 1 | 3 | 1.06 | 1.17 | 1.2 | 1.1471 | 1.19 | 1.23 | 1.18 | 1.17 |  | 16
| 1 | 4 | 1.1 | 1.16 | 1.21 | 1.1723 | 1.175 | 1.3 | 1.16 | 1.14 |  | 1.1733
| 1 | 5 | 1.05 | 1.15 | 1.19 | 1.1124 | 1.176 | 1.23 | 1.17 | 1.13 |  | 123
| 1 | 6 | 1.06 | 1.17 | 1.22 | 1.1698 | 1.186 | 1.32 | 1.15 | 1.13
| Mean |  | 1.0633 | 1.1883 | 1.225 | 1.1597 | 1.1778 | 1.2533 | 1.1633 | 1.1417
| Median |  | 1.06 | 1.17 | 1.215 | 1.1584 | 1.1755 | 1.24 | 1.16 | 1.135
| Std Dev. |  | 0.018619 | 0.044907 | 0.032711 | 0.038891 | 0.0085654 | 0.048442 | 0.010328 | 0.019408
| Rel.Std.Dev. |  | 1.751% | 3.779% | 2.67% | 3.353% | 0.7272% | 3.865% | 0.8878% | 1.7%
| PDM3 |  | -9.369% | 1.285% | 4.41% | -1.155% | 0.3902% | 6.825% | -0.8457% | -2.692%
| Z-Score (Absolute) |  | 3.1 | 0.51 | 1.57 | 0.31 | 0.21 | 2.39 | 0.21 | 0.84
| Analytical results for La in OREAS 24b (Indicative Value 29.2 ppm) | [10 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
1
| 1 | 1 | 20 | 20.4 | 38.4 | 30 | 25.92 | 39.4698 | 34.9 | 30 | 22 | 1
| 1 | 2 | 20 | 20.3 | 38 | 30 | 27.1 | 38.9717 | 35.7 | 30 | 22 | 25
| 1 | 3 | 20 | 20.4 | 37.1 | 30 | 26.97 | 39.952 | 35.4 | 30 | 23 | 16
| 1 | 4 | 30 | 19.3 | 39 | 30 | 28.72 | 39.3285 | 35.1 | 30 | 22 | 29.1621
| 1 | 5 | 20 | 19.8 | 37.4 | 30 | 26.79 | 39.0512 | 34.9 | 30 | 23 | 31
| 1 | 6 | 20 | 20.8 | 38.5 | 30 | 27.66 | 39.2212 | 36.8 | 30 | 23
| Mean |  | 21.6667 | 20.1667 | 38.0667 | 30 | 27.1933 | 39.3324 | 35.4667 | 30 | 22.5
| Median |  | 20 | 20.35 | 38.2 | 30 | 27.035 | 39.2748 | 35.25 | 30 | 22.5
| Std Dev. |  | 4.0825 | 0.53166 | 0.71461 | 0 | 0.93714 | 0.35336 | 0.72296 | 0 | 0.54772
| Rel.Std.Dev. |  | 18.84% | 2.636% | 1.877% | 0% | 3.446% | 0.8984% | 2.038% | 0% | 2.434%
| PDM3 |  | -25.7% | -30.85% | 30.53% | 2.873% | -6.751% | 34.87% | 21.62% | 2.873% | -22.85%
| Z-Score (Absolute) |  | 0.75 | 0.88 | 0.73 | 0 | 0.25 | 0.84 | 0.49 | 0 | 0.67
| Analytical results for Li in OREAS 24b (Certified Value 45.6 ppm) | [10 empty] | Certified
Li | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 07 | ppm
|  |  | 01: AR*MS; 02: 2A*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS | 1
1
| 1 | 1 | 01: 40.1; 02: 40; 03: 48.3; 04: 43.25; 05: 61 | 1
| 1 | 2 | 01: 39.2; 02: 45; 03: 49.7; 04: 43.1338; 05: 59.1 | 0
| 1 | 3 | 01: 38.2; 02: 42; 03: 46.7; 04: 43.0049; 05: 54.1 | 16
| 1 | 4 | 01: 38.7; 02: 43; 03: 49.4; 04: 42.5736; 05: 56.1 | 45.6248
| 1 | 5 | 01: 39.4; 02: 42; 03: 44.7; 04: 43.0386; 05: 50.4 | 124
| 1 | 6 | 01: 38.5; 02: 43; 03: 47.2; 04: 42.8436; 05: 55.1
| Mean |  | 01: 39.0167; 02: 42.5; 03: 47.6667; 04: 42.9741; 05: 55.9667
| Median |  | 01: 38.95; 02: 42.5; 03: 47.75; 04: 43.0218; 05: 55.6
| Std Dev. |  | 01: 0.69113; 02: 1.6432; 03: 1.8705; 04: 0.23838; 05: 3.7522
| Rel.Std.Dev. |  | 01: 1.771%; 02: 3.866%; 03: 3.924%; 04: 0.5547%; 05: 6.704%
| PDM3 |  | 01: -14.48%; 02: -6.849%; 03: 4.475%; 04: -5.81%; 05: 22.67%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.08; 03: 0.8; 04: 0; 05: 2.21
| Analytical results for Lu in OREAS 24b (Indicative Value 0.2 ppm) | [10 empty] | Indicative
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab | 1
| No | No | 01: 02; 02: 04; 03: 06; 04: 07 | ppm
|  |  | 01: AR*MS; 02: 2A*OES/MS; 03: AR*OES/MS; 04: AR*MS | 2
2
| 1 | 1 | 01: 0.13; 02: 0.2; 03: 0.19052; 04: 0.2 | 1
| 1 | 2 | 01: 0.13; 02: 0.2; 03: 0.19043; 04: 0.2 | 17
| 1 | 3 | 01: 0.12; 02: 0.2; 03: 0.18824; 04: 0.2 | 16
| 1 | 4 | 01: 0.125; 02: 0.2; 03: 0.18508; 04: 0.2 | 0.19566
| 1 | 5 | 01: 0.125; 02: 0.2; 03: 0.18243; 04: 0.2 | 32
| 1 | 6 | 01: 0.125; 02: 0.2; 03: 0.18519; 04: 0.2
| Mean |  | 01: 0.12583; 02: 0.2; 03: 0.18698; 04: 0.2
| Median |  | 01: 0.125; 02: 0.2; 03: 0.18671; 04: 0.2
| Std Dev. |  | 01: 0.0037639; 02: 3.0405e-17; 03: 0.0032729; 04: 3.0405e-17
| Rel.Std.Dev. |  | 01: 2.991%; 02: 1.52e-14%; 03: 1.75%; 04: 1.52e-14%
| PDM3 |  | 01: -35.69%; 02: 2.218%; 03: -4.436%; 04: 2.218%
| Z-Score (Absolute) |  | 01: 7.01; 02: 0.67; 03: 0.67; 04: 0.67
| Analytical results for Mg in OREAS 24b (Certified Value 1.36 wt.%) | [10 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
3
| 1 | 1 | 1.32 | 1.46 | 1.4 | 1.29 | 1.3652 | 1.52 | 1.3 | 1.28 |  | 1
| 1 | 2 | 1.3 | 1.42 | 1.44 | 1.4 | 1.3727 | 1.53 | 1.3 | 1.31 |  | 0
| 1 | 3 | 1.32 | 1.37 | 1.38 | 1.31 | 1.3905 | 1.53 | 1.32 | 1.31 |  | 16
| 1 | 4 | 1.36 | 1.43 | 1.41 | 1.34 | 1.3802 | 1.44 | 1.3 | 1.27 |  | 1.3646
| 1 | 5 | 1.3 | 1.39 | 1.36 | 1.26 | 1.3709 | 1.49 | 1.3 | 1.29 |  | 125
| 1 | 6 | 1.3 | 1.4 | 1.4 | 1.34 | 1.3902 | 1.47 | 1.29 | 1.28
| Mean |  | 1.3167 | 1.4117 | 1.3983 | 1.3233 | 1.3783 | 1.4967 | 1.3017 | 1.29
| Median |  | 1.31 | 1.41 | 1.4 | 1.325 | 1.3764 | 1.505 | 1.3 | 1.285
| Std Dev. |  | 0.023381 | 0.031885 | 0.027142 | 0.048442 | 0.010503 | 0.036697 | 0.0098319 | 0.016733
| Rel.Std.Dev. |  | 1.776% | 2.259% | 1.941% | 3.661% | 0.762% | 2.452% | 0.7553% | 1.297%
| PDM3 |  | -3.511% | 3.451% | 2.474% | -3.023% | 1.005% | 9.68% | -4.61% | -5.465%
| Z-Score (Absolute) |  | 0.48 | 0.85 | 0.66 | 0.38 | 0.38 | 2.03 | 0.69 | 0.85
| Analytical results for Mn in OREAS 24b (Certified Value 0.035 wt.%) | [10 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 3
3
| 1 | 1 | 0.0329 | 0.0379 | 0.0366 | 0.0327 | 0.035168 | 0.0423 | 0.0354 | 0.034 |  | 1
| 1 | 2 | 0.0323 | 0.0372 | 0.0374 | 0.0355 | 0.035225 | 0.0411 | 0.0351 | 0.0355 |  | 0
| 1 | 3 | 0.0328 | 0.0361 | 0.0366 | 0.0332 | 0.035734 | 0.0419 | 0.0358 | 0.0349 |  | 16
| 1 | 4 | 0.0339 | 0.0373 | 0.0363 | 0.034 | 0.035325 | 0.0407 | 0.0353 | 0.0338 |  | 0.034992
| 1 | 5 | 0.0324 | 0.0361 | 0.0354 | 0.032 | 0.035174 | 0.041 | 0.0352 | 0.0345 |  | 126
| 1 | 6 | 0.0323 | 0.0369 | 0.0369 | 0.0341 | 0.035739 | 0.0402 | 0.0348 | 0.0342
| Mean |  | 0.032767 | 0.036917 | 0.036533 | 0.033583 | 0.035394 | 0.0412 | 0.035267 | 0.034483
| Median |  | 0.0326 | 0.03705 | 0.0366 | 0.0336 | 0.035275 | 0.04105 | 0.03525 | 0.03435
| Std Dev. |  | 0.0006121 | 0.0007111 | 0.00066833 | 0.0012287 | 0.00027113 | 0.0007746 | 0.00033267 | 0.00063061
| Rel.Std.Dev. |  | 1.868% | 1.926% | 1.829% | 3.659% | 0.766% | 1.88% | 0.9433% | 1.829%
| PDM3 |  | -6.36% | 5.5% | 4.405% | -4.026% | 1.149% | 17.74% | 0.7849% | -1.454%
| Z-Score (Absolute) |  | 1.24 | 0.77 | 0.58 | 0.84 | 0.03 | 2.84 | 0.03 | 0.41
| Analytical results for Mo in OREAS 24b (Certified Value 3.86 ppm) | [10 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 2
3
| 1 | 1 | 4 | 4 | 4.2 | 3.4 | 3.5 | 3.8952 | 3.87 | 4 | 3.3 | 1
| 1 | 2 | 4 | 4 | 4.2 | 3.6 | 3.8 | 3.9648 | 3.78 | 4 | 3.5 | 9
| 1 | 3 | 4 | 4 | 4.1 | 3.5 | 3.5 | 3.8255 | 3.78 | 4 | 3.4 | 16
| 1 | 4 | 4 | 3.9 | 4.3 | 3.6 | 3.6 | 3.8043 | 3.73 | 4 | 3.4 | 3.8606
| 1 | 5 | 4 | 3.9 | 4.1 | 3.5 | 3.5 | 3.8427 | 3.64 | 4 | 3.2 | 127
| 1 | 6 | 4 | 3.9 | 4.2 | 3.6 | 3.7 | 4.1708 | 3.71 | 4 | 3.4
| Mean |  | 4 | 3.95 | 4.1833 | 3.5333 | 3.6 | 3.9172 | 3.7517 | 4 | 3.3667
| Median |  | 4 | 3.95 | 4.2 | 3.55 | 3.55 | 3.8689 | 3.755 | 4 | 3.4
| Std Dev. |  | 0 | 0.054772 | 0.075277 | 0.08165 | 0.12649 | 0.13695 | 0.077825 | 0 | 0.10328
| Rel.Std.Dev. |  | 0% | 1.387% | 1.799% | 2.311% | 3.514% | 3.496% | 2.074% | 0% | 3.068%
| PDM3 |  | 3.611% | 2.316% | 8.36% | -8.477% | -6.75% | 1.466% | -2.822% | 3.611% | -12.79%
| Z-Score (Absolute) |  | 0.34 | 0.13 | 1.08 | 1.56 | 1.29 | 0 | 0.67 | 0.34 | 2.24
| Analytical results for Na in OREAS 24b (Certified Value 0.108 wt.%) | [10 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 3
3
| 1 | 1 | 0.1 | 0.11 | 0.1 | 0.12 | 0.10791 | 0.155 | 0.11 | 0.097 |  | 1
| 1 | 2 | 0.1 | 0.1 | 0.12 | 0.13 | 0.10719 | 0.153 | 0.11 | 0.1 |  | 0
| 1 | 3 | 0.11 | 0.1 | 0.1 | 0.13 | 0.10944 | 0.153 | 0.11 | 0.102 |  | 16
| 1 | 4 | 0.11 | 0.1 | 0.11 | 0.13 | 0.1085 | 0.146 | 0.11 | 0.1 |  | 0.10783
| 1 | 5 | 0.1 | 0.1 | 0.1 | 0.12 | 0.10798 | 0.142 | 0.11 | 0.101 |  | 128
| 1 | 6 | 0.1 | 0.1 | 0.11 | 0.13 | 0.10799 | 0.142 | 0.11 | 0.1
| Mean |  | 0.10333 | 0.10167 | 0.10667 | 0.12667 | 0.10817 | 0.1485 | 0.11 | 0.1
| Median |  | 0.1 | 0.1 | 0.105 | 0.13 | 0.10799 | 0.1495 | 0.11 | 0.1
| Std Dev. |  | 0.005164 | 0.0040825 | 0.008165 | 0.005164 | 0.00075111 | 0.0058907 | 0 | 0.0016733
| Rel.Std.Dev. |  | 4.997% | 4.016% | 7.655% | 4.077% | 0.6944% | 3.967% | 0% | 1.673%
| PDM3 |  | -4.173% | -5.719% | -1.082% | 17.46% | 0.3113% | 37.71% | 2.009% | -7.265%
| Z-Score (Absolute) |  | 0.56 | 0.79 | 0.1 | 2.64 | 0.1 | 5.63 | 0.35 | 1.02
| Analytical results for Nb in OREAS 24b (Indicative Value 0.31 ppm) | [10 empty] | Indicative
Nb | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 01: 02; 02: 05; 04: 06; 05: 07 |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*OES/MS; 05: AR*MS |  | 2
2
| 1 | 1 | 01: < 0.2; 02: < 0.2; 04: 0.82661; 05: 0.3 |  | 1
| 1 | 2 | 01: < 0.2; 02: 0.5; 04: 0.79492; 05: 0.3 |  | 27
| 1 | 3 | 01: < 0.2; 02: 0.3; 04: 0.78824; 05: 0.3 |  | 16
| 1 | 4 | 01: < 0.2; 02: 0.3; 04: 0.74516; 05: 0.3 |  | 0.30833
| 1 | 5 | 01: < 0.2; 02: 0.4; 04: 0.74517; 05: 0.4 |  | 33
| 1 | 6 | 01: < 0.2; 02: 0.3; 04: 0.80144; 05: 0.3
| Mean |  | 02: 0.36; 04: 0.78359; 05: 0.31667
| Median |  | 02: 0.3; 04: 0.79158; 05: 0.3
| Std Dev. |  | 02: 0.089443; 04: 0.032474; 05: 0.040825
| Rel.Std.Dev. |  | 02: 24.85%; 04: 4.144%; 05: 12.89%
| PDM3 |  | 02: 16.76%; 04: 154.1%; 05: 2.703%
| Z-Score (Absolute) |  | 01: 1.35; 02: 0; 04: 2.91; 05: 0
| Analytical results for Nd in OREAS 24b (Indicative Value 24.6 ppm) | [10 empty] | Indicative
Nd | Round | Replicate | 01: Lab; 02: Lab |  | 1
| No | No | 01: 02; 02: 07 |  | ppm
|  |  | 01: AR*MS; 02: AR*MS |  | 1
1
| 1 | 1 | 01: 19.9; 02: 30.7 |  | 1
| 1 | 2 | 01: 18.9; 02: 30.7 |  | 18
| 1 | 3 | 01: 18.5; 02: 30.9 |  | 16
| 1 | 4 | 01: 17.7; 02: 29.4 |  | 24.6333
| 1 | 5 | 01: 18.4; 02: 30.6 |  | 34
| 1 | 6 | 01: 19.3; 02: 30.6
| Mean |  | 01: 18.7833; 02: 30.4833
| Median |  | 01: 18.7; 02: 30.65
| Std Dev. |  | 01: 0.76529; 02: 0.54191
| Rel.Std.Dev. |  | 01: 4.074%; 02: 1.778%
| PDM3 |  | 01: -23.75%; 02: 23.75%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Ni in OREAS 24b (Certified Value 57 ppm) | [10 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*OES | AR*MS | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | 0
1
| 1 | 1 | 50 | 58 | 60 | 60 | 51.5 | 51.746 | 62.9 | 54 | 58.3 | 1
| 1 | 2 | 49 | 56 | 58 | 60 | 56.8 | 52.502 | 64.6 | 55 | 60.9 | 3
| 1 | 3 | 50 | 55 | 61 | 50 | 52.5 | 52.9765 | 59.9 | 55 | 59.9 | 16
| 1 | 4 | 53 | 56 | 60 | 60 | 54.1 | 53.5698 | 65.1 | 55 | 56.1 | 56.5306
| 1 | 5 | 50 | 55 | 59 | 60 | 50.8 | 55.1902 | 59.6 | 55 | 55.7 | 129
| 1 | 6 | 51 | 56 | 59 | 60 | 53.5 | 54.6667 | 67.4 | 57 | 58.4
| Mean |  | 50.5 | 56 | 59.5 | 58.3333 | 53.2 | 53.4418 | 63.25 | 55.1667 | 58.2167
| Median |  | 50 | 56 | 59.5 | 60 | 53 | 53.2731 | 63.75 | 55 | 58.35
| Std Dev. |  | 1.3784 | 1.0954 | 1.0488 | 4.0825 | 2.1448 | 1.3076 | 3.0703 | 0.98319 | 2.0439
| Rel.Std.Dev. |  | 2.73% | 1.956% | 1.763% | 6.999% | 4.032% | 2.447% | 4.854% | 1.782% | 3.511%
| PDM3 |  | -10.67% | -0.9386% | 5.253% | 3.189% | -5.892% | -5.464% | 11.89% | -2.413% | 2.983%
| Z-Score (Absolute) |  | 1.45 | 0 | 0.92 | 0.62 | 0.74 | 0.67 | 1.91 | 0.22 | 0.58
| Analytical results for P in OREAS 24b (Indicative Value 0.062 wt.%) | [10 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 |  |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 3
3
| 1 | 1 | 0.058 | 0.064 | 0.062 | 0.0513 | 0.062083 | 0.062 | 0.054 |  |  | 1
| 1 | 2 | 0.057 | 0.064 | 0.062 | 0.0554 | 0.062008 | 0.062 | 0.059 |  |  | 10
| 1 | 3 | 0.058 | 0.06 | 0.063 | 0.0525 | 0.062702 | 0.063 | 0.06 |  |  | 16
| 1 | 4 | 0.06 | 0.064 | 0.063 | 0.0549 | 0.062597 | 0.062 | 0.056 |  |  | 0.062229
| 1 | 5 | 0.057 | 0.06 | 0.062 | 0.0509 | 0.062803 | 0.062 | 0.057 |  |  | 35
| 1 | 6 | 0.057 | 0.062 | 0.061 | 0.0527 | 0.06231 | 0.061 | 0.059
| Mean |  | 0.057833 | 0.062333 | 0.062167 | 0.05295 | 0.062417 | 0.062 | 0.0575
| Median |  | 0.0575 | 0.063 | 0.062 | 0.0526 | 0.062453 | 0.062 | 0.058
| Std Dev. |  | 0.001169 | 0.0019664 | 0.00075277 | 0.0018436 | 0.00033268 | 0.00063246 | 0.0022583
| Rel.Std.Dev. |  | 2.021% | 3.155% | 1.211% | 3.482% | 0.533% | 1.02% | 3.928%
| PDM3 |  | -7.064% | 0.1673% | -0.1006% | -14.91% | 0.3017% | -0.3684% | -7.6%
| Z-Score (Absolute) |  | 6.74 | 0.54 | 0.27 | 14.63 | 0.67 | 0 | 7.28
| Analytical results for Pb in OREAS 24b (Certified Value 9.23 ppm) | [10 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 2
2
| 1 | 1 | 9 | 9 | 9.7 | 10 | 8.5 | 7.6 | 11.2 | 9 | 8.8 | 1
| 1 | 2 | 8 | 9 | 9.6 | 10 | 9.4 | 7.85 | 10.8 | 9 | 9 | 11
| 1 | 3 | 8 | 11 | 9.4 | 9 | 8.3 | 8.46 | 11.4 | 9 | 8.8 | 16
| 1 | 4 | 9 | 11 | 9.8 | 10 | 8.9 | 8.26 | 10.2 | 9 | 8.5 | 9.2339
| 1 | 5 | 9 | 12 | 9.6 | 9 | 8.3 | 8.21 | 11 | 8 | 10.1 | 130
| 1 | 6 | 8 | 9 | 9.5 | 10 | 8.8 | 8.87 | 10.3 | 8 | 8.8
| Mean |  | 8.5 | 10.1667 | 9.6 | 9.6667 | 8.7 | 8.2083 | 10.8167 | 8.6667 | 9
| Median |  | 8.5 | 10 | 9.6 | 10 | 8.65 | 8.235 | 10.9 | 9 | 8.8
| Std Dev. |  | 0.54772 | 1.3292 | 0.14142 | 0.5164 | 0.42426 | 0.44772 | 0.48339 | 0.5164 | 0.56214
| Rel.Std.Dev. |  | 6.444% | 13.07% | 1.473% | 5.342% | 4.877% | 5.455% | 4.469% | 5.958% | 6.246%
| PDM3 |  | -7.948% | 10.1% | 3.965% | 4.687% | -5.782% | -11.11% | 17.14% | -6.143% | -2.533%
| Z-Score (Absolute) |  | 0.56 | 1.31 | 0.67 | 0.75 | 0.34 | 0.89 | 2.04 | 0.37 | 0
| Analytical results for Pd in OREAS 24b (Indicative Value < 0.01 ppm) | [10 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 05 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 3
3
| 1 | 1 | 01: < 0.01; 02: < 0.01 | 1
| 1 | 2 | 01: < 0.01; 02: < 0.01 | 1
| 1 | 3 | 01: < 0.01; 02: < 0.01 | 16
| 1 | 4 | 01: < 0.01; 02: < 0.01 | < 0.01
| 1 | 5 | 01: < 0.01; 02: < 0.01 | 36
| 1 | 6 | 01: < 0.01; 02: < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 02
| Analytical results for Pr in OREAS 24b (Indicative Value 6.87 ppm) | [10 empty] | Indicative
Pr | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 02; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 2
2
| 1 | 1 | 01: 5.14; 02: 8.9 | 1
| 1 | 2 | 01: 4.98; 02: 8.8 | 19
| 1 | 3 | 01: 4.87; 02: 8.8 | 16
| 1 | 4 | 01: 4.63; 02: 8.6 | 6.8658
| 1 | 5 | 01: 4.99; 02: 8.8 | 37
| 1 | 6 | 01: 4.98; 02: 8.9
| Mean |  | 01: 4.9317; 02: 8.8
| Median |  | 01: 4.98; 02: 8.8
| Std Dev. |  | 01: 0.17105; 02: 0.10954
| Rel.Std.Dev. |  | 01: 3.468%; 02: 1.245%
| PDM3 |  | 01: -28.17%; 02: 28.17%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Pt in OREAS 24b (Indicative Value 0.001 ppm) | [10 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 03; 02: 05 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 3
3
| 1 | 1 | 01: 0.001; 02: < 0.005 | 1
| 1 | 2 | 01: 0.001; 02: < 0.005 | 1
| 1 | 3 | 01: 0.001; 02: < 0.005 | 16
| 1 | 4 | 01: 0.002; 02: < 0.005 | 0.0013333
| 1 | 5 | 01: 0.001; 02: < 0.005 | 38
| 1 | 6 | 01: 0.002; 02: < 0.005
| Mean |  | 01: 0.0013333
| Median |  | 01: 0.001
| Std Dev. |  | 01: 0.0005164
| Rel.Std.Dev. |  | 01: 38.73%
| PDM3 |  | 01: 2.442e-13%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Rb in OREAS 24b (Certified Value 114 ppm) | [10 empty] | Certified
Rb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 07 | ppm
|  |  | 01: AR*MS; 02: 2A*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS | 0
0
| 1 | 1 | 01: 104; 02: 109; 03: 103.86; 04: 119.0373; 05: 136 | 1
| 1 | 2 | 01: 102; 02: 112; 03: 109.96; 04: 116.8887; 05: 139 | 0
| 1 | 3 | 01: 106; 02: 111; 03: 104.05; 04: 113.4167; 05: 127 | 16
| 1 | 4 | 01: 110; 02: 115; 03: 110.85; 04: 114.3314; 05: 131 | 114.3427
| 1 | 5 | 01: 106; 02: 112; 03: 101.53; 04: 112.9942; 05: 121 | 131
| 1 | 6 | 01: 107; 02: 113; 03: 108.05; 04: 116.3128; 05: 138
| Mean |  | 01: 105.8333; 02: 112; 03: 106.3833; 04: 115.4968; 05: 132
| Median |  | 01: 106; 02: 112; 03: 106.05; 04: 115.3221; 05: 133.5
| Std Dev. |  | 01: 2.7142; 02: 2; 03: 3.7653; 04: 2.3275; 05: 7.0427
| Rel.Std.Dev. |  | 01: 2.565%; 02: 1.786%; 03: 3.539%; 04: 2.015%; 05: 5.335%
| PDM3 |  | 01: -7.442%; 02: -2.049%; 03: -6.961%; 04: 1.009%; 05: 15.44%
| Z-Score (Absolute) |  | 01: 0.74; 02: 0; 03: 0.67; 04: 0.42; 05: 2.4
| Analytical results for Re in OREAS 24b (Indicative Value < 0.001 ppm) | [10 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 05; 02: 07 | ppm
|  |  | 01: AR*MS; 02: AR*MS | 3
3
| 1 | 1 | 01: < 0.05; 02: < 0.001 | 1
| 1 | 2 | 01: < 0.05; 02: < 0.001 | 12
| 1 | 3 | 01: < 0.05; 02: < 0.001 | 16
| 1 | 4 | 01: < 0.05; 02: < 0.001 | < 0.001
| 1 | 5 | 01: < 0.05; 02: < 0.001 | 39
| 1 | 6 | 01: < 0.05; 02: < 0.001
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for S in OREAS 24b (Certified Value 0.2 wt.%) | [10 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 06 | 08 | 09 |  |  |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 3
3
| 1 | 1 | 0.19 | 0.21 | 0.22 | 0.19671 | 0.21 | 0.18 |  |  |  | 1
| 1 | 2 | 0.19 | 0.205 | 0.21 | 0.1968 | 0.21 | 0.19 |  |  |  | 4
| 1 | 3 | 0.19 | 0.2 | 0.21 | 0.20218 | 0.21 | 0.19 |  |  |  | 16
| 1 | 4 | 0.2 | 0.21 | 0.21 | 0.19888 | 0.21 | 0.19 |  |  |  | 0.19992
| 1 | 5 | 0.19 | 0.2 | 0.2 | 0.20021 | 0.21 | 0.18 |  |  |  | 132
| 1 | 6 | 0.19 | 0.205 | 0.21 | 0.20238 | 0.2 | 0.19
| Mean |  | 0.19167 | 0.205 | 0.21 | 0.19953 | 0.20833 | 0.18667
| Median |  | 0.19 | 0.205 | 0.21 | 0.19954 | 0.21 | 0.19
| Std Dev. |  | 0.0040825 | 0.0044721 | 0.0063246 | 0.002506 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 2.13% | 2.182% | 3.012% | 1.256% | 1.96% | 2.766%
| PDM3 |  | -4.129% | 2.541% | 5.042% | -0.1979% | 4.208% | -6.63%
| Z-Score (Absolute) |  | 1.03 | 0.27 | 0.76 | 0.27 | 0.59 | 1.52
| Analytical results for Sb in OREAS 24b (Indicative Value 0.48 ppm) | [10 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
2
| 1 | 1 | < 2 | 0.48 | 0.3 | 0.55 | 0.42641 | 0.61 | < 2 | 0.4 |  | 1
| 1 | 2 | < 2 | 0.54 | 0.3 | 0.7 | 0.4209 | 0.63 | < 2 | 0.3 |  | 7
| 1 | 3 | < 2 | 0.52 | 0.3 | 0.57 | 0.45686 | 0.63 | < 2 | 0.4 |  | 16
| 1 | 4 | < 2 | 0.5 | 0.3 | 0.55 | 0.43895 | 0.6 | < 2 | 0.4 |  | 0.47841
| 1 | 5 | < 2 | 0.5 | 0.3 | 0.6 | 0.42375 | 0.65 | < 2 | 0.4 |  | 28
| 1 | 6 | < 2 | 0.54 | 0.3 | 0.61 | 0.46605 | 0.61 | < 2 | 0.4
| Mean |  |  | 0.51333 | 0.3 | 0.59667 | 0.43882 | 0.62167 |  | 0.38333
| Median |  |  | 0.51 | 0.3 | 0.585 | 0.43268 | 0.62 |  | 0.4
| Std Dev. |  |  | 0.024221 | 0 | 0.056451 | 0.018812 | 0.018348 |  | 0.040825
| Rel.Std.Dev. |  |  | 4.718% | 0% | 9.461% | 4.287% | 2.951% |  | 10.65%
| PDM3 |  |  | 7.299% | -37.29% | 24.72% | -8.276% | 29.94% |  | -19.87%
| Z-Score (Absolute) |  | 2.08 | 0.2 | 1.19 | 0.2 | 0.54 | 0.31 | 2.08 | 0.8
| Analytical results for Sc in OREAS 24b (Certified Value 9.51 ppm) | [10 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
3
| 1 | 1 | 9 | 10 | 9.9 | 9 | 8.12 | 10.9 | 10 | 8.6 |  | 1
| 1 | 2 | 9 | 9 | 10.6 | 10 | 8.37 | 10.9 | 10 | 9.5 |  | 0
| 1 | 3 | 9 | 9 | 10 | 9 | 8.78 | 10.1 | 10 | 9 |  | 16
| 1 | 4 | 9 | 10 | 10.6 | 10 | 8.77 | 10.3 | 10 | 8.8 |  | 9.5128
| 1 | 5 | 9 | 9 | 10 | 9 | 8.65 | 10 | 10 | 8.9 |  | 133
| 1 | 6 | 9 | 9 | 10.2 | 9 | 8.69 | 11.3 | 10 | 9.1
| Mean |  | 9 | 9.3333 | 10.2167 | 9.3333 | 8.5633 | 10.5833 | 10 | 8.9833
| Median |  | 9 | 9 | 10.1 | 9 | 8.67 | 10.6 | 10 | 8.95
| Std Dev. |  | 0 | 0.5164 | 0.31252 | 0.5164 | 0.26349 | 0.52313 | 0 | 0.30605
| Rel.Std.Dev. |  | 0% | 5.533% | 3.059% | 5.533% | 3.077% | 4.943% | 0% | 3.407%
| PDM3 |  | -5.39% | -1.886% | 7.4% | -1.886% | -9.98% | 11.25% | 5.122% | -5.565%
| Z-Score (Absolute) |  | 0.44 | 0 | 1.17 | 0 | 1.02 | 1.66 | 0.88 | 0.46
| Analytical results for Se in OREAS 24b (Indicative Value 0.42 ppm) | [10 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 09 |  |  |  | ppm
|  |  | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 2
2
| 1 | 1 | < 1 | < 5 | < 1 | 0.54536 | 0.7 | < 0.5 |  |  |  | 1
| 1 | 2 | < 1 | < 5 | < 1 | 0.42773 | 0.6 | < 0.5 |  |  |  | 8
| 1 | 3 | < 1 | < 5 | < 1 | 0.27549 | 0.5 | < 0.5 |  |  |  | 16
| 1 | 4 | < 1 | < 5 | < 1 | 0.47771 | 0.6 | < 0.5 |  |  |  | 0.4192
| 1 | 5 | < 1 | < 5 | < 1 | 0.34073 | 0.6 | < 0.5 |  |  |  | 29
| 1 | 6 | < 1 | < 5 | < 1 | 0.3786 | 0.6 | 0.7
| Mean |  | 05: 0.40761; 06: 0.6; 07: 0.7
| Median |  | 05: 0.40317; 06: 0.6; 07: 0.7
| Std Dev. |  | 05: 0.097021; 06: 0.063246
| Rel.Std.Dev. |  | 05: 23.8%; 06: 10.54%
| PDM3 |  | 05: -2.766%; 06: 43.13%; 07: 66.98%
| Z-Score (Absolute) |  | 0 | 14.02 | 0 | 0.65 | 0.7 | 1.23
| Analytical results for Sm in OREAS 24b (Indicative Value 4.68 ppm) | [10 empty] | Indicative
Sm | Round | Replicate | 01: Lab; 02: Lab |  | 1
| No | No | 01: 02; 02: 07 |  | ppm
|  |  | 01: AR*MS; 02: AR*MS |  | 2
2
| 1 | 1 | 01: 3.73; 02: 5.9 |  | 1
| 1 | 2 | 01: 3.64; 02: 5.7 |  | 15
| 1 | 3 | 01: 3.6; 02: 5.9 |  | 16
| 1 | 4 | 01: 3.46; 02: 5.5 |  | 4.6808
| 1 | 5 | 01: 3.56; 02: 5.8 |  | 30
| 1 | 6 | 01: 3.78; 02: 5.6
| Mean |  | 01: 3.6283; 02: 5.7333
| Median |  | 01: 3.62; 02: 5.75
| Std Dev. |  | 01: 0.116; 02: 0.1633
| Rel.Std.Dev. |  | 01: 3.197%; 02: 2.848%
| PDM3 |  | 01: -22.49%; 02: 22.49%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Sn in OREAS 24b (Certified Value 2.26 ppm) | [10 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 |  |  |  | ppm
|  |  | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 2
3
| 1 | 1 | 2.3 | 2 | 2 | 2.3 | 2.376 | 2.56 |  |  |  | 1
| 1 | 2 | 2.3 | 2 | 2 | 2.5 | 2.3555 | 2.58 |  |  |  | 0
| 1 | 3 | 2.5 | 2 | 1.9 | 2.3 | 2.3941 | 2.53 |  |  |  | 16
| 1 | 4 | 2.3 | 2 | 2.1 | 2.5 | 2.4574 | 2.44 |  |  |  | 2.2602
| 1 | 5 | 2.3 | 2 | 2 | 2.3 | 2.4083 | 2.39 |  |  |  | 134
| 1 | 6 | 2.2 | 2 | 2 | 2.4 | 2.4249 | 2.47
| Mean |  | 2.3167 | 2 | 2 | 2.3833 | 2.4027 | 2.495
| Median |  | 2.3 | 2 | 2 | 2.35 | 2.4012 | 2.5
| Std Dev. |  | 0.098319 | 0 | 0.063246 | 0.098319 | 0.036136 | 0.073959
| Rel.Std.Dev. |  | 4.244% | 0% | 3.162% | 4.125% | 1.504% | 2.964%
| PDM3 |  | 2.5% | -11.51% | -11.51% | 5.449% | 6.306% | 10.39%
| Z-Score (Absolute) |  | 0.23 | 2.39 | 2.39 | 0.23 | 0.36 | 0.99
| Analytical results for Sr in OREAS 24b (Indicative Value 29 ppm) | [10 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
1
| 1 | 1 | 29 | 28.8 | 36 | 32.7 | 29.7 | 29.0181 | 42.8 | 29 | 26 | 1
| 1 | 2 | 29 | 27.8 | 36 | 38.7 | 30.6 | 28.7891 | 44.4 | 29 | 28 | 5
| 1 | 3 | 30 | 28.6 | 35 | 31.7 | 30.5 | 29.2961 | 39.7 | 30 | 28 | 16
| 1 | 4 | 31 | 28.2 | 37 | 34.8 | 32.3 | 29.0494 | 40.8 | 30 | 26 | 29.0079
| 1 | 5 | 30 | 29.2 | 35 | 34.4 | 29.1 | 28.8378 | 38.6 | 30 | 27 | 31
| 1 | 6 | 29 | 29.2 | 36 | 31.9 | 30.4 | 28.893 | 43.8 | 29 | 26
| Mean |  | 29.6667 | 28.6333 | 35.8333 | 34.0333 | 30.4333 | 28.9806 | 41.6833 | 29.5 | 26.8333
| Median |  | 29.5 | 28.7 | 36 | 33.55 | 30.45 | 28.9556 | 41.8 | 29.5 | 26.5
| Std Dev. |  | 0.8165 | 0.55737 | 0.75277 | 2.6181 | 1.0801 | 0.18461 | 2.3379 | 0.54772 | 0.98319
| Rel.Std.Dev. |  | 2.752% | 1.947% | 2.101% | 7.693% | 3.549% | 0.637% | 5.609% | 1.857% | 3.664%
| PDM3 |  | 2.271% | -1.291% | 23.53% | 17.32% | 4.914% | -0.09406% | 43.7% | 1.697% | -7.496%
| Z-Score (Absolute) |  | 0 | 0.67 | 4.02 | 2.85 | 0.5 | 0.45 | 7.84 | 0.11 | 1.85
| Analytical results for Ta in OREAS 24b (Indicative Value < 0.05 ppm) | [10 empty] | Indicative
Ta | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 06; 05: 07 | ppm
|  |  | 01: AR*MS; 02: 2A*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS | 3
3
| 1 | 1 | 01: < 0.005; 02: < 0.1; 03: < 0.05; 04: 0.0080645; 05: < 0.05 | 1
| 1 | 2 | 01: < 0.005; 02: 0.2; 03: < 0.05; 04: 0.0068359; 05: < 0.05 | 26
| 1 | 3 | 01: < 0.005; 02: < 0.1; 03: < 0.05; 04: 0.0078431; 05: < 0.05 | 16
| 1 | 4 | 01: < 0.005; 02: < 0.1; 03: < 0.05; 04: 0.0087209; 05: < 0.05 | < 0.05
| 1 | 5 | 01: < 0.005; 02: < 0.1; 03: < 0.05; 04: 0.0067568; 05: < 0.05 | 32
| 1 | 6 | 01: < 0.005; 02: < 0.1; 03: < 0.05; 04: 0.0072016; 05: < 0.05
| Mean |  | 02: 0.2; 04: 0.0075705
| Median |  | 02: 0.2; 04: 0.0075224
| Std Dev. |  | 04: 0.00077197
| Rel.Std.Dev. |  | 04: 10.2%
PDM3
| Z-Score (Absolute) |  | 01: 0.87; 02: 1.93; 03: 0; 04: 0.67; 05: 0
| Analytical results for Tb in OREAS 24b (Indicative Value 0.54 ppm) | [10 empty] | Indicative
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab | 1
| No | No | 01: 02; 02: 04; 03: 06; 04: 07 | ppm
|  |  | 01: AR*MS; 02: 2A*OES/MS; 03: AR*OES/MS; 04: AR*MS | 2
2
| 1 | 1 | 01: 0.455; 02: 0.4; 03: 0.72581; 04: 0.6 | 1
| 1 | 2 | 01: 0.435; 02: 0.4; 03: 0.71973; 04: 0.6 | 16
| 1 | 3 | 01: 0.43; 02: 0.3; 03: 0.71373; 04: 0.6 | 16
| 1 | 4 | 01: 0.43; 02: 0.4; 03: 0.70736; 04: 0.6 | 0.53775
| 1 | 5 | 01: 0.44; 02: 0.4; 03: 0.70077; 04: 0.6 | 33
| 1 | 6 | 01: 0.45; 02: 0.4; 03: 0.69856; 04: 0.6
| Mean |  | 01: 0.44; 02: 0.38333; 03: 0.71099; 04: 0.6
| Median |  | 01: 0.4375; 02: 0.4; 03: 0.71054; 04: 0.6
| Std Dev. |  | 01: 0.010488; 02: 0.040825; 03: 0.010728; 04: 0
| Rel.Std.Dev. |  | 01: 2.384%; 02: 10.65%; 03: 1.509%; 04: 0%
| PDM3 |  | 01: -18.18%; 02: -28.72%; 03: 32.22%; 04: 11.58%
| Z-Score (Absolute) |  | 01: 0.5; 02: 0.85; 03: 1.19; 04: 0.5
| Analytical results for Te in OREAS 24b (Indicative Value < 0.02 ppm) | [10 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 09 |  |  |  | ppm
|  |  | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 3
3
| 1 | 1 | < 0.02 | < 2 | < 0.05 | 0.029234 | < 0.02 | < 0.2 |  |  |  | 1
| 1 | 2 | < 0.02 | < 2 | < 0.05 | 0.023438 | < 0.02 | < 0.2 |  |  |  | 9
| 1 | 3 | < 0.02 | < 2 | < 0.05 | 0.019608 | < 0.02 | < 0.2 |  |  |  | 16
| 1 | 4 | < 0.02 | < 2 | < 0.05 | 0.027132 | < 0.02 | < 0.2 |  |  |  | < 0.02
| 1 | 5 | < 0.02 | < 2 | < 0.05 | 0.023166 | < 0.02 | < 0.2 |  |  |  | 34
| 1 | 6 | < 0.02 | < 2 | < 0.05 | 0.023663 | < 0.02 | < 0.2
| Mean |  | 04: 0.024373
| Median |  | 04: 0.02355
| Std Dev. |  | 04: 0.0033699
| Rel.Std.Dev. |  | 04: 13.83%
PDM3
| Z-Score (Absolute) |  | 0.67 | 44.78 | 0.01 | 0.01 | 0.67 | 3.46
| Analytical results for Th in OREAS 24b (Certified Value 14.3 ppm) | [10 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | 1
1
| 1 | 1 | < 20 | 10.9 | 15.1 | 12.9 | 12.5 | 15.5373 | 15.3 | < 20 | 11.3 | 1
| 1 | 2 | < 20 | 10.5 | 15 | 12.6 | 13.73 | 15.6201 | 15.4 | < 20 | 11.7 | 0
| 1 | 3 | < 20 | 10.2 | 14.8 | 12 | 12.84 | 16.8598 | 16.7 | < 20 | 11.7 | 16
| 1 | 4 | 20 | 10.3 | 15.6 | 13.1 | 13.39 | 16.2364 | 15.8 | < 20 | 11.1 | 14.3325
| 1 | 5 | < 20 | 10.5 | 15 | 12.4 | 12.61 | 16.3475 | 16 | < 20 | 11.6 | 135
| 1 | 6 | < 20 | 10.6 | 15.1 | 12.5 | 13.21 | 16.2767 | 15.8 | < 20 | 11.9
| Mean |  | 20 | 10.5 | 15.1 | 12.5833 | 13.0467 | 16.1463 | 15.8333 |  | 11.55
| Median |  | 20 | 10.5 | 15.05 | 12.55 | 13.025 | 16.2566 | 15.8 |  | 11.65
| Std Dev. |  |  | 0.24495 | 0.26833 | 0.38687 | 0.47827 | 0.49446 | 0.50067 |  | 0.29496
| Rel.Std.Dev. |  |  | 2.333% | 1.777% | 3.074% | 3.666% | 3.062% | 3.162% |  | 2.554%
| PDM3 |  | 39.54% | -26.74% | 5.355% | -12.2% | -8.971% | 12.66% | 10.47% |  | -19.41%
| Z-Score (Absolute) |  | 0.3 | 0.67 | 0.81 | 0 | 0.15 | 1.15 | 1.05 | 0.84 | 0.33
| Analytical results for Ti in OREAS 24b (Certified Value 0.198 wt.%) | [10 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 08 | 09 |  |  | wt.%
|  |  | AR*OES | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS |  |  | 3
3
| 1 | 1 | 0.2 | 0.17 | 0.19 | > 0.1638 | 0.21201 | 0.2 | 0.189 |  |  | 1
| 1 | 2 | 0.2 | 0.16 | 0.2 | > 0.1795 | 0.21196 | 0.2 | 0.193 |  |  | 0
| 1 | 3 | 0.2 | 0.16 | 0.18 | > 0.1655 | 0.21188 | 0.21 | 0.192 |  |  | 16
| 1 | 4 | 0.21 | 0.16 | 0.18 | > 0.1689 | 0.21431 | 0.2 | 0.182 |  |  | 0.19849
| 1 | 5 | 0.2 | 0.155 | 0.18 | > 0.1674 | 0.20847 | 0.2 | 0.189 |  |  | 136
| 1 | 6 | 0.21 | 0.16 | 0.19 | > 0.176 | 0.21112 | 0.2 | 0.19
| Mean |  | 0.20333 | 0.16083 | 0.18667 |  | 0.21163 | 0.20167 | 0.18917
| Median |  | 0.2 | 0.16 | 0.185 |  | 0.21192 | 0.2 | 0.1895
| Std Dev. |  | 0.005164 | 0.004916 | 0.008165 |  | 0.0018827 | 0.0040825 | 0.0038687
| Rel.Std.Dev. |  | 2.54% | 3.057% | 4.374% |  | 0.8896% | 2.024% | 2.045%
| PDM3 |  | 2.439% | -18.97% | -5.958% |  | 6.617% | 1.599% | -4.698%
| Z-Score (Absolute) |  | 0.64 | 2.8 | 0.71 | NA | 1.31 | 0.51 | 0.51
| Analytical results for Tl in OREAS 24b (Certified Value 0.66 ppm) | [10 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
2
| 1 | 1 | < 10 | 0.62 | 0.7 | 0.61 | 0.70968 | 0.77 | < 10 | 0.5 |  | 1
| 1 | 2 | < 10 | 0.63 | 0.7 | 0.69 | 0.7041 | 0.75 | < 10 | 0.5 |  | 0
| 1 | 3 | < 10 | 0.61 | 0.7 | 0.64 | 0.73824 | 0.78 | < 10 | 0.6 |  | 16
| 1 | 4 | < 10 | 0.63 | 0.7 | 0.65 | 0.72674 | 0.73 | < 10 | 0.5 |  | 0.65777
| 1 | 5 | < 10 | 0.63 | 0.7 | 0.65 | 0.72876 | 0.74 | < 10 | 0.5 |  | 137
| 1 | 6 | < 10 | 0.63 | 0.7 | 0.65 | 0.72222 | 0.74 | < 10 | 0.5
| Mean |  |  | 0.625 | 0.7 | 0.64833 | 0.72162 | 0.75167 |  | 0.51667
| Median |  |  | 0.63 | 0.7 | 0.65 | 0.72448 | 0.745 |  | 0.5
| Std Dev. |  |  | 0.0083666 | 1.2162e-16 | 0.025626 | 0.012675 | 0.019408 |  | 0.040825
| Rel.Std.Dev. |  |  | 1.339% | 1.737e-14% | 3.953% | 1.756% | 2.582% |  | 7.902%
| PDM3 |  |  | -4.982% | 6.42% | -1.435% | 9.708% | 14.27% |  | -21.45%
| Z-Score (Absolute) |  | 39.01 | 0.78 | 0.1 | 0.57 | 0.1 | 0.37 | 39.01 | 1.77
| Analytical results for Tm in OREAS 24b (Indicative Value 0.17 ppm) | [10 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab |  | 1
| No | No | 01: 02; 02: 07 |  | ppm
|  |  | 01: AR*MS; 02: AR*MS |  | 2
2
| 1 | 1 | 01: 0.15; 02: 0.2 |  | 1
| 1 | 2 | 01: 0.145; 02: 0.2 |  | 18
| 1 | 3 | 01: 0.145; 02: 0.2 |  | 16
| 1 | 4 | 01: 0.145; 02: 0.2 |  | 0.17292
| 1 | 5 | 01: 0.145; 02: 0.2 |  | 35
| 1 | 6 | 01: 0.145; 02: 0.2
| Mean |  | 01: 0.14583; 02: 0.2
| Median |  | 01: 0.145; 02: 0.2
| Std Dev. |  | 01: 0.0020412; 02: 3.0405e-17
| Rel.Std.Dev. |  | 01: 1.4%; 02: 1.52e-14%
| PDM3 |  | 01: -15.66%; 02: 15.66%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for U in OREAS 24b (Certified Value 1.74 ppm) | [10 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES |  | 2
2
| 1 | 1 | < 10 | 1.46 | 2 | 1.41 | 1.66 | 2.004 | 2 | < 10 |  | 1
| 1 | 2 | < 10 | 1.4 | 2 | 1.28 | 1.77 | 2.006 | 2 | < 10 |  | 0
| 1 | 3 | < 10 | 1.36 | 1.9 | 1.18 | 1.66 | 2.173 | 2.1 | < 10 |  | 16
| 1 | 4 | < 10 | 1.34 | 2 | 1.29 | 1.8 | 2.091 | 2 | < 10 |  | 1.7417
| 1 | 5 | < 10 | 1.38 | 2 | 1.21 | 1.65 | 2.085 | 2 | < 10 |  | 138
| 1 | 6 | < 10 | 1.42 | 2 | 1.29 | 1.71 | 2.071 | 2 | < 10
| Mean |  |  | 1.3933 | 1.9833 | 1.2767 | 1.7083 | 2.0717 | 2.0167
| Median |  |  | 1.39 | 2 | 1.285 | 1.685 | 2.078 | 2
| Std Dev. |  |  | 0.043205 | 0.040825 | 0.079917 | 0.063692 | 0.062787 | 0.040825
| Rel.Std.Dev. |  |  | 3.101% | 2.058% | 6.26% | 3.728% | 3.031% | 2.024%
| PDM3 |  |  | -20% | 13.88% | -26.7% | -1.914% | 18.95% | 15.79%
| Z-Score (Absolute) |  | 4.5 | 0.91 | 0.03 | 1.09 | 0.44 | 0.11 | 0.03 | 4.5
| Analytical results for V in OREAS 24b (Certified Value 79 ppm) | [10 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 08 | 09 |  |  | ppm
|  |  | AR*OES | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS |  |  | 0
1
| 1 | 1 | 78 | 72 | 77 | 80.7369 | 88 | 80 | 77 |  |  | 1
| 1 | 2 | 77 | 65 | 84 | 78.4951 | 83 | 80 | 78 |  |  | 0
| 1 | 3 | 78 | 71 | 77 | 80.4 | 90 | 81 | 79 |  |  | 16
| 1 | 4 | 80 | 71 | 80 | 80.1085 | 85 | 80 | 77 |  |  | 78.9213
| 1 | 5 | 78 | 69 | 76 | 79.6052 | 90 | 80 | 77 |  |  | 139
| 1 | 6 | 79 | 71 | 79 | 80.2922 | 97 | 79 | 77
| Mean |  | 78.3333 | 69.8333 | 78.8333 | 79.9397 | 88.8333 | 80 | 77.5
| Median |  | 78 | 71 | 78 | 80.2004 | 89 | 80 | 77
| Std Dev. |  | 1.0328 | 2.5626 | 2.9269 | 0.79986 | 4.8751 | 0.63246 | 0.83666
| Rel.Std.Dev. |  | 1.318% | 3.67% | 3.713% | 1.001% | 5.488% | 0.7906% | 1.08%
| PDM3 |  | -0.745% | -11.52% | -0.1114% | 1.29% | 12.56% | 1.367% | -1.801%
| Z-Score (Absolute) |  | 0.29 | 5.2 | 0 | 0.64 | 5.78 | 0.67 | 0.77
| Analytical results for W in OREAS 24b (Indicative Value 1.19 ppm) | [10 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS |  | 2
2
| 1 | 1 | < 10 | 0.9 | 1.7 | 0.9 | 1.255 | 1.3 | < 10 | 0.9 |  | 1
| 1 | 2 | < 10 | 0.8 | 1.8 | 1.2 | 1.3828 | 1 | < 10 | 1 |  | 30
| 1 | 3 | < 10 | 0.8 | 1.6 | 0.9 | 1.4902 | 1.2 | < 10 | 0.9 |  | 16
| 1 | 4 | < 10 | 0.8 | 1.7 | 0.9 | 2.14 | 1.1 | < 10 | 0.9 |  | 1.1886
| 1 | 5 | < 10 | 0.9 | 1.6 | 1 | 1.8021 | 1.3 | < 10 | 1.1 |  | 36
| 1 | 6 | < 10 | 0.8 | 1.7 | 0.9 | 1.2994 | 1.1 | < 10 | 1
| Mean |  |  | 0.83333 | 1.6833 | 0.96667 | 1.5616 | 1.1667 |  | 0.96667
| Median |  |  | 0.8 | 1.7 | 0.9 | 1.4365 | 1.15 |  | 0.95
| Std Dev. |  |  | 0.05164 | 0.075277 | 0.12111 | 0.34415 | 0.12111 |  | 0.08165
| Rel.Std.Dev. |  |  | 6.197% | 4.472% | 12.53% | 22.04% | 10.38% |  | 8.447%
| PDM3 |  |  | -29.89% | 41.62% | -18.67% | 31.38% | -1.845% |  | -18.67%
| Z-Score (Absolute) |  | 6.17 | 0.9 | 0.54 | 0.67 | 0.34 | 0.34 | 6.17 | 0.67
| Analytical results for Y in OREAS 24b (Certified Value 12.3 ppm) | [10 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 |  |  |  | ppm
|  |  | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
1
| 1 | 1 | 11.4 | 13.1 | 9.79 | 10.58 | 14.047 | 15.6 |  |  |  | 1
| 1 | 2 | 10.6 | 13.1 | 10.7 | 11.08 | 13.738 | 15.9 |  |  |  | 0
| 1 | 3 | 11.2 | 12.9 | 10.1 | 10.75 | 13.405 | 15 |  |  |  | 16
| 1 | 4 | 10.9 | 13.4 | 10.6 | 11.3 | 13.417 | 14.9 |  |  |  | 12.3449
| 1 | 5 | 10.9 | 12.8 | 9.94 | 10.56 | 13.427 | 14.7 |  |  |  | 140
| 1 | 6 | 11 | 13 | 10.5 | 10.97 | 13.711 | 15.4
| Mean |  | 11 | 13.05 | 10.2717 | 10.8733 | 13.6242 | 15.25
| Median |  | 10.95 | 13.05 | 10.3 | 10.86 | 13.569 | 15.2
| Std Dev. |  | 0.27568 | 0.20736 | 0.37812 | 0.29446 | 0.25656 | 0.45935
| Rel.Std.Dev. |  | 2.506% | 1.589% | 3.681% | 2.708% | 1.883% | 3.012%
| PDM3 |  | -10.89% | 5.712% | -16.79% | -11.92% | 10.36% | 23.53%
| Z-Score (Absolute) |  | 0.5 | 0.5 | 0.86 | 0.56 | 0.78 | 1.58
| Analytical results for Yb in OREAS 24b (Indicative Value 1.15 ppm) | [10 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 06 | 07 | [5 empty] | ppm
|  |  | AR*MS | 2A*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
2
| 1 | 1 | 0.89 | 1.5 | 1.1925 | 1.1 | [5 empty] | 1
| 1 | 2 | 0.89 | 1.5 | 1.1846 | 1.1 | [5 empty] | 19
| 1 | 3 | 0.86 | 1.4 | 1.1627 | 1.1 | [5 empty] | 16
| 1 | 4 | 0.88 | 1.5 | 1.1647 | 1 | [5 empty] | 1.1505
| 1 | 5 | 0.88 | 1.5 | 1.1486 | 1 | [5 empty] | 37
| 1 | 6 | 0.9 | 1.5 | 1.1595 | 1
| Mean |  | 0.88333 | 1.4833 | 1.1688 | 1.05
| Median |  | 0.885 | 1.5 | 1.1637 | 1.05
| Std Dev. |  | 0.013663 | 0.040825 | 0.016485 | 0.054772
| Rel.Std.Dev. |  | 1.547% | 2.752% | 1.41% | 5.216%
| PDM3 |  | -23.22% | 28.93% | 1.587% | -8.738%
| Z-Score (Absolute) |  | 1.07 | 1.77 | 0.28 | 0.28
| Analytical results for Zn in OREAS 24b (Certified Value 93 ppm) | [10 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | ppm
|  |  | AR*OES | AR*OES | AR*MS | 2A*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | 0
1
| 1 | 1 | 91 | 97 | 105.6 | 104 | 84 | 91.4829 | 115 | 96 | 92 | 1
| 1 | 2 | 91 | 94 | 104.3 | 102 | 93 | 92.8633 | 117 | 94 | 94 | 4
| 1 | 3 | 91 | 92 | 103.5 | 96 | 84 | 93.6549 | 110 | 95 | 93 | 16
| 1 | 4 | 95 | 94 | 107 | 95 | 87 | 93.5271 | 107 | 94 | 90 | 92.7511
| 1 | 5 | 91 | 93 | 104.1 | 94 | 82 | 93.6689 | 104 | 96 | 90 | 141
| 1 | 6 | 92 | 92 | 103.7 | 96 | 88 | 95.1471 | 112 | 94 | 93
| Mean |  | 91.8333 | 93.6667 | 104.7 | 97.8333 | 86.3333 | 93.3907 | 110.8333 | 94.8333 | 92
| Median |  | 91 | 93.5 | 104.2 | 96 | 85.5 | 93.591 | 111 | 94.5 | 92.5
| Std Dev. |  | 1.6021 | 1.8619 | 1.3461 | 4.1191 | 3.9328 | 1.1974 | 4.8751 | 0.98319 | 1.6733
| Rel.Std.Dev. |  | 1.745% | 1.988% | 1.286% | 4.21% | 4.555% | 1.282% | 4.399% | 1.037% | 1.819%
| PDM3 |  | -0.9894% | 0.9872% | 12.88% | 5.479% | -6.919% | 0.6896% | 19.5% | 2.245% | -0.8098%
| Z-Score (Absolute) |  | 0.67 | 0 | 4.06 | 1.53 | 2.7 | 0.1 | 6.31 | 0.43 | 0.61
| Analytical results for Zr in OREAS 24b (Certified Value 24.5 ppm) | [10 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 |  |  |  | ppm
|  |  | AR*MS | AR*MS | 2A*OES/MS | AR*MS | AR*OES/MS | AR*MS |  |  |  | 1
1
| 1 | 1 | 22 | 2.4 | 21.8 | 22 | 31.45 | 23.8 |  |  |  | 1
| 1 | 2 | 20 | 2.2 | 24.3 | 23.6 | 29.77 | 26.3 |  |  |  | 0
| 1 | 3 | 21 | 2.1 | 22.2 | 23 | 30.7 | 25.1 |  |  |  | 16
| 1 | 4 | 21 | 2.2 | 22.5 | 22.1 | 30.69 | 24.7 |  |  |  | 24.467
| 1 | 5 | 21 | 1.9 | 21.8 | 21.9 | 29.88 | 28.2 |  |  |  | 142
| 1 | 6 | 22 | 2.4 | 22.4 | 22.8 | 29.98 | 28.2
| Mean |  | 21.1667 | 2.2 | 22.5 | 22.5667 | 30.4117 | 26.05
| Median |  | 21 | 2.2 | 22.3 | 22.45 | 30.335 | 25.7
| Std Dev. |  | 0.75277 | 0.18974 | 0.92952 | 0.67725 | 0.65107 | 1.8491
| Rel.Std.Dev. |  | 3.556% | 8.624% | 4.131% | 3.001% | 2.141% | 7.098%
| PDM3 |  | -13.49% | -91.01% | -8.039% | -7.767% | 24.3% | 6.47%
| Z-Score (Absolute) |  | 0.38 | 5.62 | 0.01 | 0.01 | 2.18 | 0.97
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1151: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 24b (Certified Value 0.189 wt.%) | [10 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.19 | 0.15 | 0.19 | 0.19 | 0.2 | 0.17959 | 0.2 | 0.15 | 0.17 | 1
| 1 | 2 | 0.19 | 0.16 | 0.19 | 0.18 | 0.21 | 0.17895 | 0.19 | 0.15 | 0.17 | 0
| 1 | 3 | 0.19 | 0.15 | 0.19 | 0.19 | 0.2 | 0.18591 | 0.2 | 0.16 | 0.17 | 16
| 1 | 4 | 0.19 | 0.15 | 0.2 | 0.19 | 0.21 | 0.18201 | 0.2 | 0.16 | 0.17 | 0.18876
| 1 | 5 | 0.19 | 0.16 | 0.2 | 0.19 | 0.19 | 0.183 | 0.18 | 0.15 | 0.17 | 144
| 1 | 6 | 0.22 | 0.15 | 0.19 | 0.2 | 0.21 | 0.17833 | 0.19 | 0.16 | 0.17
| Mean |  | 0.195 | 0.15333 | 0.19333 | 0.19 | 0.20333 | 0.1813 | 0.19333 | 0.155 | 0.17
| Median |  | 0.19 | 0.15 | 0.19 | 0.19 | 0.205 | 0.1808 | 0.195 | 0.155 | 0.17
| Std Dev. |  | 0.012247 | 0.005164 | 0.005164 | 0.0063246 | 0.008165 | 0.0028953 | 0.008165 | 0.0054772 | 0
| Rel.Std.Dev. |  | 6.281% | 3.368% | 2.671% | 3.329% | 4.016% | 1.597% | 4.223% | 3.534% | 0%
| PDM3 |  | 3.307% | -18.77% | 2.425% | 0.6586% | 7.722% | -3.951% | 2.425% | -17.88% | -9.937%
| Z-Score (Absolute) |  | 0.39 | 2.84 | 0.26 | 0 | 1.03 | 0.67 | 0.26 | 2.71 | 1.55
| Analytical results for S in OREAS 24b (Certified Value 0.19 wt.%) | [10 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.18 | 0.19 | 0.17 | 0.21 | 0.192 | 0.185 | 0.19 | 0.19 | 0.15 | 1
| 1 | 2 | 0.18 | 0.2 | 0.15 | 0.21 | 0.195 | 0.186 | 0.19 | 0.19 | 0.15 | 4
| 1 | 3 | 0.19 | 0.19 | 0.16 | 0.21 | 0.202 | 0.184 | 0.19 | 0.19 | 0.16 | 16
| 1 | 4 | 0.18 | 0.2 | 0.17 | 0.23 | 0.203 | 0.185 | 0.2 | 0.2 | 0.16 | 0.19006
| 1 | 5 | 0.18 | 0.19 | 0.16 | 0.21 | 0.201 | 0.191 | 0.2 | 0.2 | 0.19 | 145
| 1 | 6 | 0.18 | 0.18 | 0.19 | 0.21 | 0.201 | 0.188 | 0.19 | 0.19 | 0.17
| Mean |  | 0.18167 | 0.19167 | 0.16667 | 0.21333 | 0.199 | 0.1865 | 0.19333 | 0.19333 | 0.16333
| Median |  | 0.18 | 0.19 | 0.165 | 0.21 | 0.201 | 0.1855 | 0.19 | 0.19 | 0.16
| Std Dev. |  | 0.0040825 | 0.0075277 | 0.013663 | 0.008165 | 0.0044272 | 0.0025884 | 0.005164 | 0.005164 | 0.015055
| Rel.Std.Dev. |  | 2.247% | 3.928% | 8.198% | 3.827% | 2.225% | 1.388% | 2.671% | 2.671% | 9.218%
| PDM3 |  | -4.417% | 0.844% | -12.31% | 12.24% | 4.702% | -1.874% | 1.721% | 1.721% | -14.06%
| Z-Score (Absolute) |  | 0.92 | 0 | 2.3 | 1.99 | 0.67 | 0.48 | 0.15 | 0.15 | 2.61
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.