## Abbreviations
Table 1. Abbreviations used for OREAS 605b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with AAS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with either AAS or ICP-OES finish as appropriate
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
5A*MS | 5-acid (HF-HNO3-HClO4-HCl-H3BO3) digest with Mass Spectrometry: ICP-MS finish
5A*OES | 5-acid (HF-HNO3-HClO4-HCl-H3BO3) digest with Optical Emmision Spectrometry [aka: A(tomic)ES, ICP-OES] finish
ABL*MS | Laser Ablation with Mass Spectrometry: ICP-MS finish
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/AAS | ore grade aqua regia digest with either AAS or ICP-OES finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
at 105°C*Oven | moisture at 105°C with drying oven finish
BF*XRF | lithium borate fusion with XRF finish
FA*AAS | fire assay with AAS finish
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
LOI*TGA | loss on ignition with Thermal Gravimetric Analyser finish
PF*OES | sodium peroxide fusion with ICP-OES finish
Text Values:
N.A. | Not Applicable (Lab 03)
N.A. | Not Applicable (Lab 22)
NSS | No Sufficient Sample (Lab 07)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 605b
Lab Slot | Laboratory Name and Location
AGAT Laboratories, Mississauga, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Reno, Nevada, USA
ALS, Santiago, Santiago Metropolitan Region, Chile
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Hermosillo, Sonora, Mexico
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
PT SGS Indo Assay Laboratories, Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
Shiva Analyticals Ltd, Bangalore North, Karnataka, India

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 605b
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Ag, Silver (ppm) | 975.1468 | 43.9913 | 952.281 | 998.0127 | 957.1221 | 993.1715
Au, Gold (ppm) | 1.7181 | 0.066102 | 1.6929 | 1.7434 | 1.7125 | 1.7238
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.583 | 0.078614 | 1.5205 | 1.6455 | 1.5773 | 1.5887
Infrared Combustion
S, Sulphur (wt.%) | 10.2943 | 0.33927 | 10.1352 | 10.4533 | 10.1918 | 10.3967
4-Acid Digestion
Ag, Silver (ppm) | 1015.032 | 24.4899 | 1003.9339 | 1026.1301 | 998.7848 | 1031.2792
Al, Aluminium (wt.%) | 6.2562 | 0.23347 | 6.1491 | 6.3634 | 6.1688 | 6.3436
As, Arsenic (ppm) | 2453.897 | 121.6588 | 2397.2081 | 2510.5858 | 2379.3183 | 2528.4756
Be, Beryllium (ppm) | 0.71014 | 0.080884 | 0.67425 | 0.74604 | 0.65252 | 0.76777
Bi, Bismuth (ppm) | 49.5801 | 4.4342 | 47.8007 | 51.3596 | 47.8499 | 51.3103
Ca, Calcium (wt.%) | 0.4235 | 0.018925 | 0.41564 | 0.43135 | 0.41334 | 0.43366
Cd, Cadmium (ppm) | 13.965 | 1.0881 | 13.46 | 14.4701 | 13.5426 | 14.3874
Ce, Cerium (ppm) | 21.2325 | 6.3688 | 17.7096 | 24.7554 | 19.9942 | 22.4708
Co, Cobalt (ppm) | 19.3561 | 0.60372 | 19.11 | 19.6023 | 18.848 | 19.8643
Cr, Chromium (ppm) | 30.3356 | 4.3552 | 28.2832 | 32.3881 | 29.2278 | 31.4435
Cs, Cesium (ppm) | 2.3102 | 0.11541 | 2.2499 | 2.3706 | 2.224 | 2.3964
Cu, Copper (wt.%) | 5.0263 | 0.10857 | 4.984 | 5.0685 | 4.9436 | 5.1089
Dy, Dysprosium (ppm) | 1.3442 | 0.15081 | 1.1928 | 1.4957 | IND | IND
Er, Erbium (ppm) | 0.639 | 0.060857 | 0.59079 | 0.68721 | IND | IND
Eu, Europium (ppm) | 0.45667 | 0.075053 | 0.3744 | 0.53894 | IND | IND
Fe, Iron (wt.%) | 7.1376 | 0.28397 | 7.0195 | 7.2558 | 7.0242 | 7.2511
Ga, Gallium (ppm) | 32.7744 | 1.8626 | 31.9118 | 33.6369 | 31.7365 | 33.8122
Gd, Gadolinium (ppm) | 1.7684 | 0.22771 | 1.5639 | 1.9729 | IND | IND
Ge, Germanium (ppm) | 0.29383 | 0.04343 | 0.25718 | 0.33049 | 0.25886 | 0.3288
Hf, Hafnium (ppm) | 2.2664 | 0.1444 | 2.1984 | 2.3343 | 2.1742 | 2.3586
Ho, Holmium (ppm) | 0.22179 | 0.026591 | 0.21078 | 0.2328 | IND | IND
In, Indium (ppm) | 3.9156 | 0.2734 | 3.7821 | 4.0491 | 3.7805 | 4.0508
K, Potassium (wt.%) | 1.1958 | 0.031793 | 1.1835 | 1.2081 | 1.1735 | 1.2181
Li, Lithium (ppm) | 28.8883 | 1.692 | 28.2306 | 29.546 | 27.7298 | 30.0468
Lu, Lutetium (ppb) | 87.6056 | 16.3521 | 74.712 | 100.4992 | IND | IND
Mg, Magnesium (ppm) | 1731.1584 | 101.5993 | 1687.3247 | 1774.9921 | 1703.2328 | 1759.084
Mn, Manganese (ppm) | 109.2601 | 4.4613 | 107.4448 | 111.0755 | 106.7956 | 111.7247
Mo, Molybdenum (ppm) | 17.3099 | 0.91718 | 16.8909 | 17.729 | 16.8916 | 17.7283
Na, Sodium (wt.%) | 0.52662 | 0.023623 | 0.51598 | 0.53726 | 0.5152 | 0.53804
Nb, Niobium (ppm) | 6.7952 | 0.35562 | 6.6506 | 6.9399 | 6.524 | 7.0664
Nd, Neodymium (ppm) | 10.3811 | 2.2543 | 7.7967 | 12.9655 | 9.7844 | 10.9779
Ni, Nickel (ppm) | 21.8785 | 0.92335 | 21.5482 | 22.2089 | 21.0639 | 22.6932
P, Phosphorus (ppm) | 623.5983 | 30.5608 | 608.2298 | 638.9668 | 602.0605 | 645.1361
Pb, Lead (ppm) | 1509.614 | 39.8088 | 1491.8825 | 1527.3455 | 1470.7052 | 1548.5229
Pr, Praseodymium (ppm) | 3.005 | 0.69324 | 2.0665 | 3.9435 | 2.7918 | 3.2182
Rb, Rubidium (ppm) | 39.5401 | 2.4354 | 38.3669 | 40.7134 | 38.5571 | 40.5232
S, Sulphur (wt.%) | 9.9623 | 0.37882 | 9.7841 | 10.1405 | 9.8558 | 10.0688
Sb, Antimony (ppm) | 404.6231 | 32.4036 | 390.2123 | 419.034 | 394.5318 | 414.7144
Sc, Scandium (ppm) | 6.3135 | 0.59348 | 6.0184 | 6.6086 | 6.0716 | 6.5554
Se, Selenium (ppm) | 86.7136 | 7.3989 | 83.2293 | 90.1978 | 83.2134 | 90.2138
Sm, Samarium (ppm) | 2.1457 | 0.2478 | 1.8879 | 2.4035 | 2.0233 | 2.2681
Sn, Tin (ppm) | 6.7852 | 0.38528 | 6.5716 | 6.9987 | 6.5326 | 7.0377
Sr, Strontium (ppm) | 401.6515 | 76.7114 | 362.2217 | 441.0813 | 387.27 | 416.033
Ta, Tantalum (ppm) | 0.54771 | 0.04535 | 0.51657 | 0.57885 | 0.51714 | 0.57829
Tb, Terbium (ppm) | 0.2185 | 0.040307 | 0.18164 | 0.25536 | 0.19551 | 0.24149
Te, Tellurium (ppm) | 32.2211 | 3.346 | 30.5876 | 33.8546 | 31.2387 | 33.2035
Th, Thorium (ppm) | 4.7941 | 1.1392 | 4.1336 | 5.4547 | 4.5483 | 5.04
Ti, Titanium (wt.%) | 0.21154 | 0.010469 | 0.20724 | 0.21584 | 0.2062 | 0.21687
Tl, Thallium (ppm) | 15.5388 | 0.77615 | 15.2501 | 15.8275 | 14.9419 | 16.1356
U, Uranium (ppm) | 3.1125 | 0.21055 | 3.0088 | 3.2161 | 2.9909 | 3.2341
V, Vanadium (ppm) | 61.7898 | 2.7439 | 60.6289 | 62.9508 | 60.3086 | 63.2711
W, Tungsten (ppm) | 30.1677 | 2.3563 | 29.0746 | 31.2607 | 29.0198 | 31.3155
Y, Yttrium (ppm) | 5.9755 | 0.2729 | 5.8694 | 6.0816 | 5.7663 | 6.1846
Yb, Ytterbium (ppm) | 0.59707 | 0.058311 | 0.55306 | 0.64108 | 0.56877 | 0.62537
Zn, Zinc (wt.%) | 0.23999 | 0.0067557 | 0.23739 | 0.24259 | 0.23643 | 0.24356
Zr, Zirconium (ppm) | 76.9506 | 5.8767 | 74.1949 | 79.7064 | 74.9903 | 78.911
Aqua Regia Digestion
Ag, Silver (ppm) | 1005.338 | 18.4305 | 993.9319 | 1016.744 | 988.1375 | 1022.5385
Al, Aluminium (wt.%) | 1.1518 | 0.072673 | 1.1182 | 1.1854 | 1.1236 | 1.18
As, Arsenic (ppm) | 2377.1354 | 178.5361 | 2299.7724 | 2454.4985 | 2307.9666 | 2446.3042
Be, Beryllium (ppm) | 0.15289 | 0.025978 | 0.13924 | 0.16653 | IND | IND
Bi, Bismuth (ppm) | 50.2386 | 2.6563 | 49.4072 | 51.0701 | 48.4567 | 52.0206
Ca, Calcium (wt.%) | 0.19695 | 0.011687 | 0.1922 | 0.20171 | 0.18946 | 0.20445
Cd, Cadmium (ppm) | 13.9828 | 0.70633 | 13.5826 | 14.3829 | 13.6164 | 14.3491
Ce, Cerium (ppm) | 10.9036 | 1.1793 | 10.2016 | 11.6057 | 10.5221 | 11.2851
Co, Cobalt (ppm) | 18.4583 | 1.181 | 17.8957 | 19.0209 | 17.9202 | 18.9964
Cr, Chromium (ppm) | 23.8484 | 2.2789 | 22.7664 | 24.9305 | 22.8696 | 24.8272
Cs, Cesium (ppm) | 1.5171 | 0.083871 | 1.4745 | 1.5596 | 1.4665 | 1.5676
Cu, Copper (wt.%) | 4.9609 | 0.1116 | 4.9116 | 5.0101 | 4.8805 | 5.0412
Fe, Iron (wt.%) | 6.9945 | 0.2956 | 6.866 | 7.1231 | 6.8712 | 7.1178
Ga, Gallium (ppm) | 8.6521 | 0.66754 | 8.2948 | 9.0094 | 8.3097 | 8.9945
Hf, Hafnium (ppm) | 0.43866 | 0.053498 | 0.41073 | 0.4666 | 0.41259 | 0.46474
Hg, Mercury (ppm) | 2.1979 | 0.17231 | 2.0982 | 2.2976 | 2.102 | 2.2939
In, Indium (ppm) | 3.8013 | 0.16599 | 3.7137 | 3.8888 | 3.6782 | 3.9244
K, Potassium (wt.%) | 0.24659 | 0.0087817 | 0.24219 | 0.25099 | 0.2383 | 0.25488
Li, Lithium (ppm) | 8.6269 | 1.0229 | 8.0698 | 9.184 | 8.3279 | 8.926
Lu, Lutetium (ppb) | 36.0987 | 5.6788 | 32.558 | 39.6394 | IND | IND
Mg, Magnesium (ppm) | 1321.5401 | 116.442 | 1268.976 | 1374.1042 | 1290.3728 | 1352.7074
Mn, Manganese (ppm) | 86.9007 | 4.3033 | 85.0728 | 88.7287 | 85.1361 | 88.6653
Mo, Molybdenum (ppm) | 16.7174 | 0.99714 | 16.2698 | 17.1651 | 16.0646 | 17.3703
Na, Sodium (wt.%) | 0.050692 | 0.0075449 | 0.047372 | 0.054013 | 0.049455 | 0.051929
Nb, Niobium (ppm) | 0.75314 | 0.13875 | 0.65436 | 0.85193 | 0.71117 | 0.79512
Ni, Nickel (ppm) | 20.8224 | 1.3846 | 20.2434 | 21.4015 | 20.1132 | 21.5317
P, Phosphorus (ppm) | 187.1444 | 14.7244 | 179.9317 | 194.357 | 175.624 | 198.6647
Pb, Lead (ppm) | 894.4079 | 44.9352 | 874.5571 | 914.2587 | 875.2152 | 913.6006
Rb, Rubidium (ppm) | 17.4651 | 1.3013 | 16.7939 | 18.1362 | 16.914 | 18.0161
Re, Rhenium (ppb) | 10.5972 | 1.2796 | 9.9907 | 11.2037 | IND | IND
S, Sulphur (wt.%) | 9.2141 | 0.44294 | 8.965 | 9.4632 | 9.0371 | 9.3912
Sb, Antimony (ppm) | 323.3045 | 40.9907 | 304.6364 | 341.9727 | 315.3797 | 331.2293
Sc, Scandium (ppm) | 1.8553 | 0.24932 | 1.7235 | 1.9872 | 1.7708 | 1.9398
Se, Selenium (ppm) | 83.975 | 8.0413 | 79.7415 | 88.2086 | 80.8677 | 87.0824
Sn, Tin (ppm) | 5.8749 | 0.29029 | 5.7191 | 6.0307 | 5.6525 | 6.0973
Sr, Strontium (ppm) | 34.1703 | 5.2619 | 31.2546 | 37.086 | 32.829 | 35.5116
Te, Tellurium (ppm) | 33.3772 | 2.446 | 32.1036 | 34.6509 | 32.2415 | 34.513
Ti, Titanium (ppm) | 457.0479 | 29.5459 | 443.4462 | 470.6496 | 447.0162 | 467.0796
Tl, Thallium (ppm) | 16.1566 | 0.85845 | 15.7234 | 16.5899 | 15.6565 | 16.6567
U, Uranium (ppm) | 1.4532 | 0.098542 | 1.3981 | 1.5083 | 1.4066 | 1.4998
V, Vanadium (ppm) | 24.4072 | 1.5476 | 23.7324 | 25.0819 | 23.4791 | 25.3353
W, Tungsten (ppm) | 8.0011 | 1.0908 | 7.3882 | 8.614 | 7.7264 | 8.2758
Y, Yttrium (ppm) | 3.3679 | 0.16419 | 3.2945 | 3.4413 | 3.2623 | 3.4734
Yb, Ytterbium (ppm) | 0.26493 | 0.025864 | 0.24161 | 0.28826 | 0.24124 | 0.28862
Zn, Zinc (wt.%) | 0.23225 | 0.009835 | 0.22797 | 0.23653 | 0.22868 | 0.23581
Zr, Zirconium (ppm) | 13.2556 | 1.5084 | 12.5419 | 13.9692 | 12.8034 | 13.7077

## Indicative Values
Table 3. Indicative Values for OREAS 605b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 3.4167 | Pt | ppb | < 5
Infrared Combustion
C | wt.% | 0.055556
4-Acid Digestion
Au | ppm | 1.6367 | La | ppm | 9.4558 | Rh | ppb | 40
Ba | ppm | 413.0619 | Pd | ppb | < 50 | Ru | ppb | < 10
Hg | ppm | < 2 | Pt | ppb | < 10 | Tm | ppb | 77.2222
Ir | ppb | < 5 | Re | ppb | 11.3704
Aqua Regia Digestion
B | ppm | 12.0768 | Ho | ppm | 0.12817 | Rh | ppb | 34.8333
Ba | ppm | 16.7432 | Ir | ppb | < 5 | Ru | ppb | < 10
Dy | ppm | 0.73539 | La | ppm | 3.5755 | Sm | ppm | 1.1243
Er | ppm | 0.32094 | Nd | ppm | 5.7766 | Ta | ppm | < 0.01
Eu | ppm | 0.16206 | Pd | ppb | < 10 | Tb | ppm | 0.13487
Gd | ppm | 0.94425 | Pr | ppm | 1.6112 | Th | ppm | 2.0075
Ge | ppm | 0.5981 | Pt | ppb | < 5 | Tm | ppb | 35.8333
Borate Fusion XRF
Al2O3 | wt.% | 12.1 | MgO | wt.% | 0.345 | SiO2 | wt.% | 51.9
CaO | wt.% | 0.595 | MnO | wt.% | 0.02 | SO3 | wt.% | 26.05
Fe2O3 | wt.% | 10.265 | Na2O | wt.% | 0.81 | TiO2 | wt.% | 0.345
K2O | wt.% | 1.455 | P2O5 | wt.% | 0.1565
Thermogravimetry
H2O- | wt.% | 1.2883 | LOI1000 | wt.% | 13.3192
Laser Ablation ICP-MS
Ag | ppm | 1045 | Hf | ppm | 3.245 | Sm | ppm | 2.22
As | ppm | 2480 | Ho | ppm | 0.245 | Sn | ppm | 7.3
Ba | ppm | 12950 | In | ppm | 3.625 | Sr | ppm | 557.5
Be | ppm | 0.7 | La | ppm | 18.6 | Ta | ppm | 0.55
Bi | ppm | 53.75 | Lu | ppb | 90 | Tb | ppm | 0.245
Cd | ppm | 17.5 | Mn | ppm | 109 | Te | ppm | 34.6
Ce | ppm | 32.25 | Mo | ppm | 17.6 | Th | ppm | 6.995
Co | ppm | 19 | Nb | ppm | 6.915 | Ti | wt.% | 0.213
Cr | ppm | 35 | Nd | ppm | 12.45 | Tl | ppm | 16.8
Cs | ppm | 2.395 | Ni | ppm | 24 | Tm | ppb | 90
Cu | wt.% | 4.68 | Pb | ppm | 1560 | U | ppm | 3.07
Dy | ppm | 1.275 | Pr | ppm | 3.54 | V | ppm | 59.3
Er | ppm | 0.605 | Rb | ppm | 38.8 | W | ppm | 30
Eu | ppm | 0.385 | Re | ppb | 30 | Y | ppm | 6.64
Ga | ppm | 31 | Sb | ppm | 457 | Yb | ppm | 0.625
Gd | ppm | 1.69 | Sc | ppm | 8.2 | Zn | wt.% | 0.2305
Ge | ppm | 6.65 | Se | ppm | < 5 | Zr | ppm | 119.5

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 605b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Ag, ppm | 975.1468 | 43.9913 | 887.1643 | 1063.1294 | 843.173 | 1107.1206 | 4.511% | 9.022% | 13.53% | 926.3895 | 1023.9042
Au, ppm | 1.7181 | 0.066102 | 1.5859 | 1.8503 | 1.5198 | 1.9164 | 3.847% | 7.695% | 11.54% | 1.6322 | 1.804
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.583 | 0.078614 | 1.4258 | 1.7402 | 1.3472 | 1.8189 | 4.966% | 9.932% | 14.9% | 1.5039 | 1.6622
Infrared Combustion
S, wt.% | 10.2943 | 0.33927 | 9.6157 | 10.9728 | 9.2765 | 11.3121 | 3.296% | 6.591% | 9.887% | 9.7796 | 10.809
4-Acid Digestion
Ag, ppm | 1015.032 | 24.4899 | 966.0523 | 1064.0118 | 941.5624 | 1088.5016 | 2.413% | 4.825% | 7.238% | 964.2804 | 1065.7836
Al, wt.% | 6.2562 | 0.23347 | 5.7893 | 6.7232 | 5.5558 | 6.9567 | 3.732% | 7.464% | 11.2% | 5.9434 | 6.5691
As, ppm | 2453.897 | 121.6588 | 2210.5793 | 2697.2146 | 2088.9205 | 2818.8734 | 4.958% | 9.916% | 14.87% | 2331.2021 | 2576.5918
Be, ppm | 0.71014 | 0.080884 | 0.54837 | 0.87191 | 0.46749 | 0.95279 | 11.39% | 22.78% | 34.17% | 0.67463 | 0.74565
Bi, ppm | 49.5801 | 4.4342 | 40.7116 | 58.4486 | 36.2774 | 62.8828 | 8.944% | 17.89% | 26.83% | 47.1011 | 52.0591
Ca, wt.% | 0.4235 | 0.018925 | 0.38565 | 0.46135 | 0.36672 | 0.48027 | 4.469% | 8.938% | 13.41% | 0.40232 | 0.44467
Cd, ppm | 13.965 | 1.0881 | 11.7888 | 16.1412 | 10.7007 | 17.2293 | 7.792% | 15.58% | 23.37% | 13.2668 | 14.6633
Ce, ppm | 21.2325 | 6.3688 | 8.4948 | 33.9702 | 2.126 | 40.339 | 30% | 59.99% | 89.99% | 20.1709 | 22.2941
Co, ppm | 19.3561 | 0.60372 | 18.1487 | 20.5636 | 17.545 | 21.1673 | 3.119% | 6.238% | 9.357% | 18.3883 | 20.3239
Cr, ppm | 30.3356 | 4.3552 | 21.6252 | 39.0461 | 17.27 | 43.4013 | 14.36% | 28.71% | 43.07% | 28.8189 | 31.8524
Cs, ppm | 2.3102 | 0.11541 | 2.0794 | 2.5411 | 1.964 | 2.6565 | 4.996% | 9.991% | 14.99% | 2.1947 | 2.4257
Cu, wt.% | 5.0263 | 0.10857 | 4.8091 | 5.2434 | 4.7005 | 5.352 | 2.16% | 4.32% | 6.48% | 4.7749 | 5.2776
Dy, ppm | 1.3442 | 0.15081 | 1.0426 | 1.6459 | 0.89182 | 1.7967 | 11.22% | 22.44% | 33.66% | 1.277 | 1.4115
Er, ppm | 0.639 | 0.060857 | 0.51729 | 0.76071 | 0.45643 | 0.82157 | 9.524% | 19.05% | 28.57% | 0.60705 | 0.67095
Eu, ppm | 0.45667 | 0.075053 | 0.30656 | 0.60677 | 0.23151 | 0.68183 | 16.43% | 32.87% | 49.3% | 0.43383 | 0.4795
Fe, wt.% | 7.1376 | 0.28397 | 6.5697 | 7.7056 | 6.2857 | 7.9895 | 3.978% | 7.957% | 11.94% | 6.7808 | 7.4945
Ga, ppm | 32.7744 | 1.8626 | 29.0492 | 36.4995 | 27.1866 | 38.3621 | 5.683% | 11.37% | 17.05% | 31.1356 | 34.4131
Gd, ppm | 1.7684 | 0.22771 | 1.313 | 2.2238 | 1.0852 | 2.4515 | 12.88% | 25.75% | 38.63% | 1.68 | 1.8568
Ge, ppm | 0.29383 | 0.04343 | 0.20697 | 0.38069 | 0.16354 | 0.42412 | 14.78% | 29.56% | 44.34% | 0.27914 | 0.30852
Hf, ppm | 2.2664 | 0.1444 | 1.9776 | 2.5552 | 1.8332 | 2.6996 | 6.372% | 12.74% | 19.11% | 2.1531 | 2.3797
Ho, ppm | 0.22179 | 0.026591 | 0.16861 | 0.27497 | 0.14201 | 0.30156 | 11.99% | 23.98% | 35.97% | 0.2107 | 0.23288
In, ppm | 3.9156 | 0.2734 | 3.3688 | 4.4624 | 3.0954 | 4.7358 | 6.982% | 13.96% | 20.95% | 3.7198 | 4.1114
K, wt.% | 1.1958 | 0.031793 | 1.1322 | 1.2594 | 1.1004 | 1.2912 | 2.659% | 5.317% | 7.976% | 1.136 | 1.2556
Li, ppm | 28.8883 | 1.692 | 25.5043 | 32.2723 | 23.8123 | 33.9643 | 5.857% | 11.71% | 17.57% | 27.4439 | 30.3327
Lu, ppb | 87.6056 | 16.3521 | 54.9013 | 120.3098 | 38.5492 | 136.6619 | 18.67% | 37.33% | 56% | 83.2253 | 91.9858
Mg, ppm | 1731.1584 | 101.5993 | 1527.9598 | 1934.3571 | 1426.3605 | 2035.9564 | 5.869% | 11.74% | 17.61% | 1644.6005 | 1817.7164
Mn, ppm | 109.2601 | 4.4613 | 100.3376 | 118.1827 | 95.8764 | 122.6439 | 4.083% | 8.166% | 12.25% | 103.7971 | 114.7232
Mo, ppm | 17.3099 | 0.91718 | 15.4755 | 19.1443 | 14.5584 | 20.0615 | 5.299% | 10.6% | 15.9% | 16.4444 | 18.1754
Na, wt.% | 0.52662 | 0.023623 | 0.47937 | 0.57387 | 0.45575 | 0.59749 | 4.486% | 8.972% | 13.46% | 0.50029 | 0.55295
Nb, ppm | 6.7952 | 0.35562 | 6.084 | 7.5065 | 5.7283 | 7.8621 | 5.233% | 10.47% | 15.7% | 6.4555 | 7.135
Nd, ppm | 10.3811 | 2.2543 | 5.8725 | 14.8897 | 3.6182 | 17.1441 | 21.72% | 43.43% | 65.15% | 9.8621 | 10.9002
Ni, ppm | 21.8785 | 0.92335 | 20.0318 | 23.7252 | 19.1085 | 24.6486 | 4.22% | 8.441% | 12.66% | 20.7846 | 22.9725
P, ppm | 623.5983 | 30.5608 | 562.4768 | 684.7198 | 531.916 | 715.2806 | 4.901% | 9.801% | 14.7% | 592.4184 | 654.7782
Pb, ppm | 1509.614 | 39.8088 | 1429.9965 | 1589.2316 | 1390.1877 | 1629.0403 | 2.637% | 5.274% | 7.911% | 1434.1333 | 1585.0947
Pr, ppm | 3.005 | 0.69324 | 1.6185 | 4.3915 | 0.92529 | 5.0847 | 23.07% | 46.14% | 69.21% | 2.8548 | 3.1553
Rb, ppm | 39.5401 | 2.4354 | 34.6693 | 44.4109 | 32.234 | 46.8463 | 6.159% | 12.32% | 18.48% | 37.5631 | 41.5171
S, wt.% | 9.9623 | 0.37882 | 9.2046 | 10.7199 | 8.8258 | 11.0988 | 3.803% | 7.605% | 11.41% | 9.4642 | 10.4604
Sb, ppm | 404.6231 | 32.4036 | 339.816 | 469.4303 | 307.4124 | 501.8339 | 8.008% | 16.02% | 24.03% | 384.392 | 424.8543
Sc, ppm | 6.3135 | 0.59348 | 5.1265 | 7.5005 | 4.5331 | 8.0939 | 9.4% | 18.8% | 28.2% | 5.9978 | 6.6292
Se, ppm | 86.7136 | 7.3989 | 71.9157 | 101.5115 | 64.5168 | 108.9104 | 8.533% | 17.07% | 25.6% | 82.3779 | 91.0493
Sm, ppm | 2.1457 | 0.2478 | 1.6501 | 2.6413 | 1.4023 | 2.8891 | 11.55% | 23.1% | 34.65% | 2.0384 | 2.253
Sn, ppm | 6.7852 | 0.38528 | 6.0146 | 7.5557 | 5.6293 | 7.941 | 5.678% | 11.36% | 17.03% | 6.4459 | 7.1244
Sr, ppm | 401.6515 | 76.7114 | 248.2288 | 555.0743 | 171.5174 | 631.7856 | 19.1% | 38.2% | 57.3% | 381.5689 | 421.7341
Ta, ppm | 0.54771 | 0.04535 | 0.45701 | 0.63841 | 0.41166 | 0.68376 | 8.28% | 16.56% | 24.84% | 0.52033 | 0.5751
Tb, ppm | 0.2185 | 0.040307 | 0.13789 | 0.29911 | 0.097578 | 0.33942 | 18.45% | 36.89% | 55.34% | 0.20758 | 0.22942
Te, ppm | 32.2211 | 3.346 | 25.5291 | 38.9131 | 22.1831 | 42.2591 | 10.38% | 20.77% | 31.15% | 30.61 | 33.8321
Th, ppm | 4.7941 | 1.1392 | 2.5157 | 7.0726 | 1.3765 | 8.2118 | 23.76% | 47.53% | 71.29% | 4.5544 | 5.0338
Ti, wt.% | 0.21154 | 0.010469 | 0.1906 | 0.23247 | 0.18013 | 0.24294 | 4.949% | 9.898% | 14.85% | 0.20096 | 0.22211
Tl, ppm | 15.5388 | 0.77615 | 13.9865 | 17.0911 | 13.2103 | 17.8672 | 4.995% | 9.99% | 14.98% | 14.7618 | 16.3157
U, ppm | 3.1125 | 0.21055 | 2.6914 | 3.5336 | 2.4808 | 3.7441 | 6.765% | 13.53% | 20.29% | 2.9568 | 3.2681
V, ppm | 61.7898 | 2.7439 | 56.3021 | 67.2776 | 53.5582 | 70.0215 | 4.441% | 8.881% | 13.32% | 58.7004 | 64.8793
W, ppm | 30.1677 | 2.3563 | 25.4551 | 34.8802 | 23.0988 | 37.2365 | 7.811% | 15.62% | 23.43% | 28.6593 | 31.676
Y, ppm | 5.9755 | 0.2729 | 5.4297 | 6.5213 | 5.1568 | 6.7942 | 4.567% | 9.134% | 13.7% | 5.6767 | 6.2743
Yb, ppm | 0.59707 | 0.058311 | 0.48045 | 0.71369 | 0.42214 | 0.77201 | 9.766% | 19.53% | 29.3% | 0.56722 | 0.62692
Zn, wt.% | 0.23999 | 0.0067557 | 0.22648 | 0.2535 | 0.21973 | 0.26026 | 2.815% | 5.63% | 8.445% | 0.22799 | 0.25199
Zr, ppm | 76.9506 | 5.8767 | 65.1973 | 88.704 | 59.3207 | 94.5806 | 7.637% | 15.27% | 22.91% | 73.1031 | 80.7982
Aqua Regia Digestion
Ag, ppm | 1005.338 | 18.4305 | 968.4769 | 1042.1991 | 950.0463 | 1060.6296 | 1.833% | 3.667% | 5.5% | 955.0711 | 1055.6049
Al, wt.% | 1.1518 | 0.072673 | 1.0065 | 1.2972 | 0.93379 | 1.3698 | 6.309% | 12.62% | 18.93% | 1.0942 | 1.2094
As, ppm | 2377.1354 | 178.5361 | 2020.0632 | 2734.2076 | 1841.5271 | 2912.7438 | 7.511% | 15.02% | 22.53% | 2258.2787 | 2495.9922
Be, ppm | 0.15289 | 0.025978 | 0.10093 | 0.20484 | 0.074955 | 0.23082 | 16.99% | 33.98% | 50.97% | 0.14524 | 0.16053
Bi, ppm | 50.2386 | 2.6563 | 44.926 | 55.5513 | 42.2697 | 58.2076 | 5.287% | 10.57% | 15.86% | 47.7267 | 52.7506
Ca, wt.% | 0.19695 | 0.011687 | 0.17358 | 0.22033 | 0.16189 | 0.23201 | 5.934% | 11.87% | 17.8% | 0.18711 | 0.2068
Cd, ppm | 13.9828 | 0.70633 | 12.5701 | 15.3954 | 11.8638 | 16.1018 | 5.051% | 10.1% | 15.15% | 13.2836 | 14.6819
Ce, ppm | 10.9036 | 1.1793 | 8.5451 | 13.2621 | 7.3658 | 14.4414 | 10.82% | 21.63% | 32.45% | 10.3584 | 11.4488
Co, ppm | 18.4583 | 1.181 | 16.0963 | 20.8203 | 14.9154 | 22.0012 | 6.398% | 12.8% | 19.19% | 17.5354 | 19.3812
Cr, ppm | 23.8484 | 2.2789 | 19.2905 | 28.4063 | 17.0116 | 30.6852 | 9.556% | 19.11% | 28.67% | 22.656 | 25.0408
Cs, ppm | 1.5171 | 0.083871 | 1.3493 | 1.6848 | 1.2655 | 1.7687 | 5.528% | 11.06% | 16.59% | 1.4412 | 1.5929
Cu, wt.% | 4.9609 | 0.1116 | 4.7377 | 5.1841 | 4.6261 | 5.2957 | 2.25% | 4.499% | 6.749% | 4.7128 | 5.2089
Fe, wt.% | 6.9945 | 0.2956 | 6.4033 | 7.5857 | 6.1077 | 7.8813 | 4.226% | 8.452% | 12.68% | 6.6448 | 7.3442
Ga, ppm | 8.6521 | 0.66754 | 7.317 | 9.9872 | 6.6495 | 10.6547 | 7.715% | 15.43% | 23.15% | 8.2195 | 9.0847
Hf, ppm | 0.43866 | 0.053498 | 0.33167 | 0.54566 | 0.27817 | 0.59916 | 12.2% | 24.39% | 36.59% | 0.41673 | 0.4606
Hg, ppm | 2.1979 | 0.17231 | 1.8533 | 2.5425 | 1.681 | 2.7148 | 7.839% | 15.68% | 23.52% | 2.088 | 2.3078
In, ppm | 3.8013 | 0.16599 | 3.4693 | 4.1333 | 3.3033 | 4.2993 | 4.367% | 8.734% | 13.1% | 3.6112 | 3.9913
K, wt.% | 0.24659 | 0.0087817 | 0.22903 | 0.26415 | 0.22024 | 0.27293 | 3.561% | 7.123% | 10.68% | 0.23426 | 0.25892
Li, ppm | 8.6269 | 1.0229 | 6.5811 | 10.6728 | 5.5582 | 11.6957 | 11.86% | 23.71% | 35.57% | 8.1956 | 9.0583
Lu, ppb | 36.0987 | 5.6788 | 24.7411 | 47.4563 | 19.0622 | 53.1351 | 15.73% | 31.46% | 47.19% | 34.2937 | 37.9036
Mg, ppm | 1321.5401 | 116.442 | 1088.656 | 1554.4241 | 972.214 | 1670.8662 | 8.811% | 17.62% | 26.43% | 1255.4631 | 1387.6171
Mn, ppm | 86.9007 | 4.3033 | 78.2942 | 95.5073 | 73.9909 | 99.8106 | 4.952% | 9.904% | 14.86% | 82.5557 | 91.2458
Mo, ppm | 16.7174 | 0.99714 | 14.7232 | 18.7117 | 13.726 | 19.7088 | 5.965% | 11.93% | 17.89% | 15.8816 | 17.5533
Na, wt.% | 0.050692 | 0.0075449 | 0.035603 | 0.065782 | 0.028058 | 0.073327 | 14.88% | 29.77% | 44.65% | 0.048158 | 0.053227
Nb, ppm | 0.75314 | 0.13875 | 0.47563 | 1.0307 | 0.33688 | 1.1694 | 18.42% | 36.85% | 55.27% | 0.71549 | 0.7908
Ni, ppm | 20.8224 | 1.3846 | 18.0531 | 23.5917 | 16.6685 | 24.9764 | 6.65% | 13.3% | 19.95% | 19.7813 | 21.8636
P, ppm | 187.1444 | 14.7244 | 157.6956 | 216.5931 | 142.9712 | 231.3175 | 7.868% | 15.74% | 23.6% | 177.7871 | 196.5016
Pb, ppm | 894.4079 | 44.9352 | 804.5374 | 984.2783 | 759.6022 | 1029.2136 | 5.024% | 10.05% | 15.07% | 849.6875 | 939.1283
Rb, ppm | 17.4651 | 1.3013 | 14.8624 | 20.0677 | 13.561 | 21.3691 | 7.451% | 14.9% | 22.35% | 16.5918 | 18.3383
Re, ppb | 10.5972 | 1.2796 | 8.0379 | 13.1565 | 6.7583 | 14.4361 | 12.08% | 24.15% | 36.23% | 10.0674 | 11.1271
S, wt.% | 9.2141 | 0.44294 | 8.3283 | 10.1 | 7.8853 | 10.543 | 4.807% | 9.614% | 14.42% | 8.7534 | 9.6748
Sb, ppm | 323.3045 | 40.9907 | 241.3232 | 405.2858 | 200.3326 | 446.2765 | 12.68% | 25.36% | 38.04% | 307.1393 | 339.4697
Sc, ppm | 1.8553 | 0.24932 | 1.3567 | 2.354 | 1.1074 | 2.6033 | 13.44% | 26.88% | 40.31% | 1.7626 | 1.9481
Se, ppm | 83.975 | 8.0413 | 67.8925 | 100.0576 | 59.8512 | 108.0989 | 9.576% | 19.15% | 28.73% | 79.7763 | 88.1738
Sn, ppm | 5.8749 | 0.29029 | 5.2943 | 6.4555 | 5.004 | 6.7458 | 4.941% | 9.882% | 14.82% | 5.5811 | 6.1686
Sr, ppm | 34.1703 | 5.2619 | 23.6465 | 44.6942 | 18.3845 | 49.9561 | 15.4% | 30.8% | 46.2% | 32.4618 | 35.8788
Te, ppm | 33.3772 | 2.446 | 28.4852 | 38.2692 | 26.0393 | 40.7152 | 7.328% | 14.66% | 21.99% | 31.7084 | 35.0461
Ti, ppm | 457.0479 | 29.5459 | 397.9562 | 516.1396 | 368.4103 | 545.6855 | 6.464% | 12.93% | 19.39% | 434.1955 | 479.9003
Tl, ppm | 16.1566 | 0.85845 | 14.4397 | 17.8735 | 13.5813 | 18.732 | 5.313% | 10.63% | 15.94% | 15.3488 | 16.9645
U, ppm | 1.4532 | 0.098542 | 1.2561 | 1.6503 | 1.1576 | 1.7488 | 6.781% | 13.56% | 20.34% | 1.3805 | 1.5259
V, ppm | 24.4072 | 1.5476 | 21.312 | 27.5023 | 19.7644 | 29.0499 | 6.341% | 12.68% | 19.02% | 23.1868 | 25.6275
W, ppm | 8.0011 | 1.0908 | 5.8195 | 10.1827 | 4.7287 | 11.2735 | 13.63% | 27.27% | 40.9% | 7.601 | 8.4011
Y, ppm | 3.3679 | 0.16419 | 3.0395 | 3.6963 | 2.8753 | 3.8604 | 4.875% | 9.75% | 14.63% | 3.1995 | 3.5363
Yb, ppm | 0.26493 | 0.025864 | 0.2132 | 0.31666 | 0.18734 | 0.34253 | 9.763% | 19.53% | 29.29% | 0.25169 | 0.27818
Zn, wt.% | 0.23225 | 0.009835 | 0.21258 | 0.25192 | 0.20274 | 0.26175 | 4.235% | 8.469% | 12.7% | 0.22063 | 0.24386
Zr, ppm | 13.2556 | 1.5084 | 10.2387 | 16.2724 | 8.7302 | 17.7809 | 11.38% | 22.76% | 34.14% | 12.5928 | 13.9183

## Fire Assay
| Analytical results for Ag in OREAS 605b (Certified Value 975 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 00 | 01 | 05 | 06 | 09 | 13 | 14 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 25 | 26 | [9 empty] | ppm
|  |  | INAA | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*GRAV | FA*AAS | FA*GRAV | FA*GRAV | FA*GRAV | [9 empty] | 0
|  |  | 0.085g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | [9 empty] | 0
| 1 | 1 | 1017.4082 | 986.7457 | 1025 | 1005 | 963 | 915 | 985 | 983 | 967 | 938 | 998 | 912.83 | 1091.6 | 922.2403 | 1006.9 | 1017 | [9 empty] | 1
| 1 | 2 | 1064.08 | 997.5347 | 1005 | 994 | 972 | 926 | 979 | 970 | 964 | 963 | 1023 | 908.57 | 1027.3 | 865.2116 | 1014.21 | 1025 | [9 empty] | 25
| 1 | 3 | 1020.0388 | 984.1251 | 984 | 964 | 964 | 928 | 989 | 969 | 974 | 943 | 1036 | 861.62 | 1020.2 | 898.4836 | 1004.31 | 1039 | [9 empty] | 16
| 1 | 4 | 1042.9525 | 979.2523 | 984 | 977 | 979 | 949 | 986 | 980 | 958 | 951 | 996 | 874.53 | 1047.6 | 894.9338 | 1011.01 | 978 | [9 empty] | 975.1468
| 1 | 5 | 995.3723 | 984.4702 | 950 | 977 | 970 | 941 | 980 | 986 | 974 | 955 | 995 | 939.12 | 1087.8 | 954.273 | 1004.5 | 1011 | [9 empty] | 7
| 1 | 6 | 1070.954 | 991.3923 | 990 | 917 | 970 | 929 | 980 | 993 | 977 | 973 | 1020 | 915.6 | 1037.4 | 847.0616 | 999.99 | 993
|  | 7 | 1059.8343
|  | 8 | 1016.9591
|  | 9 | 1024.6127
|  | 10 | 1134.7487
|  | 11 | 1048.9489
|  | 12 | 1078.3025
|  | 13 | 1036.9225
|  | 14 | 1017.1207
|  | 15 | 1036.1577
|  | 16 | 1013.0269
|  | 17 | 1009.3223
|  | 18 | 1011.3548
|  | 19 | 1044.5719
|  | 20 | 1011.8383
| Mean |  | 1037.7264 | 987.2534 | 989.6667 | 972.3333 | 969.6667 | 931.3333 | 983.1667 | 980.1667 | 969 | 953.8333 | 1011.3333 | 902.045 | 1051.9833 | 897.034 | 1006.82 | 1010.5
| Median |  | 1030.3852 | 985.6079 | 987 | 977 | 970 | 928.5 | 982.5 | 981.5 | 970.5 | 953 | 1009 | 910.7 | 1042.5 | 896.7087 | 1005.7 | 1014
| Std Dev. |  | 32.347 | 6.3976 | 24.9853 | 30.6964 | 5.8195 | 11.9778 | 4.0702 | 9.3256 | 7.2664 | 12.875 | 17.3167 | 28.6651 | 30.6713 | 38.5437 | 5.1089 | 22.0341
| Rel.Std.Dev. |  | 3.117% | 0.648% | 2.525% | 3.157% | 0.6002% | 1.286% | 0.414% | 0.9514% | 0.7499% | 1.35% | 1.712% | 3.178% | 2.916% | 4.297% | 0.5074% | 2.181%
| PDM3 |  | 6.417% | 1.242% | 1.489% | -0.2885% | -0.562% | -4.493% | 0.8224% | 0.5148% | -0.6303% | -2.186% | 3.711% | -7.496% | 7.879% | -8.01% | 3.248% | 3.625%
| Z-Score (Absolute) |  | NA | 0.18 | 0.24 | 0.2 | 0.27 | 1.25 | 0.08 | 0 | 0.29 | 0.67 | 0.8 | 2 | 1.84 | 2.13 | 0.68 | 0.78
| Analytical results for Au in OREAS 605b (Certified Value 1.72 ppm) | [26 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 50g | 25g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 50g | 25g | 30g | 30g | 30g | 30g | 40g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 1.7345 | 1.7077 | 1.658 | 1.734 | 1.67 | 1.69 | 1.725 | 1.795 | 1.755 | 1.755 | 1.75 | 1.836 | 1.7 | 1.771 | 1.728 | 1.73 | 1.7 | 1.74 | 1.772 | 1.75 | 1.67 | 1.798 | 1.6947 | 1.5661 | 1.645 | 1
| 1 | 2 | 1.7968 | 1.7366 | 1.678 | 1.697 | 1.66 | 1.77 | 1.78 | NSS | 1.755 | 1.745 | 1.73 | 1.803 | 1.75 | 1.571 | 1.689 | 1.78 | 1.68 | 1.75 | 1.781 | 1.78 | 1.72 | 1.804 | 1.652 | 1.5949 | 1.605 | 5
| 1 | 3 | 1.7233 | 1.7263 | 1.593 | 1.719 | 1.65 | 1.72 | 1.765 | 1.73 | 1.745 | 1.735 | 1.72 | 1.826 | 1.75 | 1.572 | 1.67 | 1.69 | 1.74 | 1.76 | 1.802 | 1.75 | 1.7 | 1.793 | 1.75 | 1.6152 | 1.715 | 16
| 1 | 4 | 1.7407 | 1.7088 | 1.622 | 1.691 | 1.66 | 1.69 | 1.77 | 1.835 | 1.745 | 1.76 | 1.7 | 1.758 | 1.75 | 1.699 | 1.681 | 1.71 | 1.75 | 1.74 | 1.76 | 1.77 | 1.62 | 1.815 | 1.7043 | 1.5739 | 1.551 | 1.7181
| 1 | 5 | 1.7523 | 1.6738 | 1.639 | 1.667 | 1.66 | 1.725 | 1.785 | 1.83 | 1.605 | 1.785 | 1.69 | 1.743 | 1.76 | 1.656 | 1.696 | 1.74 | 1.75 | 1.75 | 1.808 | 1.76 | 1.63 | 1.793 | 1.73 | 1.6121 | 1.598 | 8
| 1 | 6 | 1.7536 | 1.7108 | 1.643 | 1.731 | 1.65 | 1.685 | 1.8 | 1.835 | 1.535 | 1.765 | 1.71 | 1.888 | 1.73 | 1.617 | 1.656 | 1.74 | 1.76 | 1.76 | 1.784 | 1.79 | 1.66 | 1.771 | 1.6247 | 1.5974 | 1.626
|  | 7 | 00: 1.735
|  | 8 | 00: 1.8349
|  | 9 | 00: 1.7117
|  | 10 | 00: 1.7809
|  | 11 | 00: 1.7908
|  | 12 | 00: 1.7944
|  | 13 | 00: 1.7471
|  | 14 | 00: 1.7582
|  | 15 | 00: 1.7671
|  | 16 | 00: 1.7259
|  | 17 | 00: 1.7814
|  | 18 | 00: 1.7089
|  | 19 | 00: 1.8078
|  | 20 | 00: 1.7898
| Mean |  | 1.7618 | 1.7107 | 1.6388 | 1.7065 | 1.6583 | 1.7133 | 1.7708 | 1.805 | 1.69 | 1.7575 | 1.7167 | 1.809 | 1.74 | 1.6477 | 1.6867 | 1.7317 | 1.73 | 1.75 | 1.7845 | 1.7667 | 1.6667 | 1.7957 | 1.6926 | 1.5933 | 1.6233
| Median |  | 1.7559 | 1.7098 | 1.641 | 1.708 | 1.66 | 1.705 | 1.775 | 1.83 | 1.745 | 1.7575 | 1.715 | 1.8145 | 1.75 | 1.6365 | 1.685 | 1.735 | 1.745 | 1.75 | 1.7825 | 1.765 | 1.665 | 1.7955 | 1.6995 | 1.5961 | 1.6155
| Std Dev. |  | 0.034428 | 0.021402 | 0.029335 | 0.02609 | 0.0075277 | 0.032506 | 0.025577 | 0.045139 | 0.095656 | 0.017248 | 0.021602 | 0.053389 | 0.021909 | 0.07805 | 0.024704 | 0.030605 | 0.032249 | 0.0089443 | 0.018042 | 0.01633 | 0.038816 | 0.014638 | 0.047115 | 0.019824 | 0.054935
| Rel.Std.Dev. |  | 1.954% | 1.251% | 1.79% | 1.529% | 0.4539% | 1.897% | 1.444% | 2.501% | 5.66% | 0.9814% | 1.258% | 2.951% | 1.259% | 4.737% | 1.465% | 1.767% | 1.864% | 0.5111% | 1.011% | 0.9243% | 2.329% | 0.8152% | 2.784% | 1.244% | 3.384%
| PDM3 |  | 2.539% | -0.4349% | -4.615% | -0.6769% | -3.48% | -0.2792% | 3.067% | 5.056% | -1.637% | 2.291% | -0.08515% | 5.289% | 1.273% | -4.101% | -1.831% | 0.7879% | 0.6909% | 1.855% | 3.863% | 2.825% | -2.995% | 4.513% | -1.485% | -7.268% | -5.517%
| Z-Score (Absolute) |  | NA | 0.06 | 1.13 | 0.13 | 0.84 | 0.02 | 0.83 | 1.34 | 0.37 | 0.63 | 0.02 | 1.4 | 0.37 | 1 | 0.42 | 0.25 | 0.22 | 0.52 | 1.03 | 0.77 | 0.72 | 1.2 | 0.33 | 1.81 | 1.36
| Analytical results for Pd in OREAS 605b (Indicative Value 3.42 ppb) | [26 empty] | Indicative
Pd | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 16; 01: 23 | ppb
|  |  | 00: FA*OES; 01: FA*AAS | 2
|  |  | 00: 40g; 01: 30g | 2
| 1 | 1 | 00: 5; 01: 3 | 1
| 1 | 2 | 00: 5; 01: < 3 | 1
| 1 | 3 | 00: 5; 01: < 3 | 16
| 1 | 4 | 00: 5; 01: < 3 | 3.4167
| 1 | 5 | 00: < 5; 01: 3 | 7
| 1 | 6 | 00: 5; 01: 3
| Mean |  | 00: 5; 01: 3
| Median |  | 00: 5; 01: 3
| Std Dev. |  | 00: 0; 01: 0
Rel.Std.Dev. | bar: 00 to 01
| PDM3 |  | 00: 46.34%; 01: -12.2%
| Z-Score (Absolute) |  | 00: 0.67; 01: 0.67
| Analytical results for Pt in OREAS 605b (Indicative Value < 5 ppb) | [26 empty] | Indicative
Pt | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 16; 01: 23 | ppb
|  |  | 00: FA*OES; 01: FA*AAS | 2
|  |  | 00: 40g; 01: 30g | 2
| 1 | 1 | 00: < 5; 01: < 5 | 1
| 1 | 2 | 00: < 5; 01: < 5 | 1
| 1 | 3 | 00: < 5; 01: < 5 | 16
| 1 | 4 | 00: 5; 01: < 5 | < 5
| 1 | 5 | 00: 5; 01: < 5 | 7
| 1 | 6 | 00: 5; 01: < 5
| Mean |  | 00: 5
| Median |  | 00: 5
| Std Dev. |  | 00: 0
Rel.Std.Dev. | bar: 00
PDM3
| Z-Score (Absolute) |  | 00: 0.67; 01: 0.67
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C63: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C81: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C99: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 605b (Certified Value 1.58 ppm) | [14 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 16 | 23 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | 2
|  |  | 0.085g | 15g | 40g | 25g | 25g | 25g | 25g | 25g | 30g | 10g | 15g | 40g | 25g | 3
| 1 | 1 | 1.7345 | N.A. | 1.58 | 1.57 | 1.62 | > 1 | 1.67 | > 1 | 1.49 | 1.85 | 1.505 | 1.67 | 1.5143 | 1
| 1 | 2 | 1.7968 | N.A. | 1.58 | 1.54 | 1.66 | > 1 | 1.76 | > 1 | 1.54 | 1.987 | 1.507 | 1.59 | 1.5751 | 5
| 1 | 3 | 1.7233 | N.A. | 1.58 | 1.54 | 1.64 | > 1 | 1.56 | > 1 | 1.52 | 1.928 | 1.482 | 1.72 | 1.5827 | 16
| 1 | 4 | 1.7407 | N.A. | 1.59 | 1.57 | 1.65 | > 1 | 1.66 | > 1 | 1.54 | 1.93 | 1.486 | 1.7 | 1.5485 | 1.583
| 1 | 5 | 1.7523 | N.A. | 1.59 | 1.71 | 1.65 | > 1 | 1.7 | > 1 | 1.51 | 1.953 | 1.488 | 1.74 | 1.5095 | 10
| 1 | 6 | 1.7536 | N.A. | 1.58 | 1.54 | 1.58 | > 1 | 1.53 | > 1 | 1.49 | 1.855 | 1.492 | 1.72 | 1.4526
|  | 7 | 1.735
|  | 8 | 1.8349
|  | 9 | 1.7117
|  | 10 | 1.7809
|  | 11 | 1.7908
|  | 12 | 1.7944
|  | 13 | 1.7471
|  | 14 | 1.7582
|  | 15 | 1.7671
|  | 16 | 1.7259
|  | 17 | 1.7814
|  | 18 | 1.7089
|  | 19 | 1.8078
|  | 20 | 1.7898
| Mean |  | 1.7618 |  | 1.5833 | 1.5783 | 1.6333 |  | 1.6467 |  | 1.515 | 1.9172 | 1.4933 | 1.69 | 1.5305
| Median |  | 1.7559 |  | 1.58 | 1.555 | 1.645 |  | 1.665 |  | 1.515 | 1.929 | 1.49 | 1.71 | 1.5314
| Std Dev. |  | 0.034428 |  | 0.005164 | 0.066156 | 0.029439 |  | 0.086641 |  | 0.022583 | 0.05444 | 0.010347 | 0.054406 | 0.048612
| Rel.Std.Dev. |  | 1.954% |  | 0.3261% | 4.192% | 1.802% |  | 5.262% |  | 1.491% | 2.84% | 0.6929% | 3.219% | 3.176%
| PDM3 |  | 11.29% |  | 0.02014% | -0.2957% | 3.179% |  | 4.021% |  | -4.297% | 21.11% | -5.665% | 6.758% | -3.321%
| Z-Score (Absolute) |  | NA | NA | 0 | 0.05 | 0.53 | NA | 0.67 | NA | 0.73 | 3.55 | 0.96 | 1.14 | 0.56
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 605b (Indicative Value 0.056 wt.%) | [20 empty] | Indicative
C | Round | Replicate | Lab | BV Geo | Lab | [16 empty] | 1
| No | No | 15 | Perth | 23 | [16 empty] | wt.%
|  |  | IRC | IRC | IRC | [16 empty] | 3
3
| 1 | 1 | 0.05 |  | 0.05 | [16 empty] | 1
| 1 | 2 | 0.05 | 0.07 | 0.05 | [16 empty] | 3
| 1 | 3 | 0.06 |  | 0.04 | [16 empty] | 16
| 1 | 4 | 0.05 | 0.07 | 0.04 | [16 empty] | 0.055556
| 1 | 5 | 0.05 |  | 0.05 | [16 empty] | 9
| 1 | 6 | 0.05 |  | 0.04
| Mean |  | 0.051667 | 0.07 | 0.045
| Median |  | 0.05 | 0.07 | 0.045
| Std Dev. |  | 0.0040825 | 0 | 0.0054772
| Rel.Std.Dev. |  | 7.902% | 0% | 12.17%
| PDM3 |  | -7% | 26% | -19%
| Z-Score (Absolute) |  | 0 | 1.85 | 0.67
| Analytical results for S in OREAS 605b (Certified Value 10.29 wt.%) | [20 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 14 | 15 | 16 | 19 | 20 | 21 | 22 | 23 | 25 | 26 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 9.84 | 10.14 | 10.8 | 9.92 | 10.5 | 10 | 10.5 | 10.7 | 10.3 | 10.7 | 10.6 | 10.4 | 10.6 | 10.6 | 10.1 | 10.3 | 9.52 | 10.0495 | 9.97 | 1
| 1 | 2 | 10.31 | 10.13 | 10.8 | 9.92 | 10.4 | 10.3 | 10.55 | 10.8 | 10.26 | 10.63 | 10.6 | 10.5 | 10.5 | 10.64 | 10 | 10.3 | 9.66 | 10.0415 | 10 | #N/A
| 1 | 3 | 10.24 | 10.12 | 10.9 | 9.74 | 10.45 | 10.5 | 10.25 | 10.8 | 10.31 | 10.68 | 10.7 | 10.4 | 10.7 | 10.55 | 9.94 | 10.3 | 9.53 | 10.0295 | 9.89 | 16
| 1 | 4 | 9.72 | 10.14 | 10.9 | 9.92 | 10.4 | 9.84 | 10.6 | 10.8 | 10.24 | 10.73 | 10.4 | 10.5 | 10.5 | 10.5 | 10.1 | 10.3 | 9.62 | 10.001 | 9.89 | 10.2943
| 1 | 5 | 10.21 | 10.15 | 10.6 | 9.74 | 10.35 | 9.85 | 10.45 | 10.7 | 10.07 | 10.73 | 10.5 | 10.4 | 10.4 | 10.66 | 10 | 10.2 | 9.61 | 10.0055 | 9.98 | 12
| 1 | 6 | 10.25 | 10.13 | 10.8 | 10.1 | 10.45 | 9.99 | 10.2 | 10.8 | 10.36 | 10.73 | 10.6 | 10.5 | 10.6 | 10.59 | 9.95 | 10.3 | 9.7 | 10.0095 | 9.87
| Mean |  | 10.095 | 10.135 | 10.8 | 9.89 | 10.425 | 10.08 | 10.425 | 10.7667 | 10.2567 | 10.7 | 10.5667 | 10.45 | 10.55 | 10.59 | 10.015 | 10.2833 | 9.6067 | 10.0228 | 9.9333
| Median |  | 10.225 | 10.135 | 10.8 | 9.92 | 10.425 | 9.995 | 10.475 | 10.8 | 10.28 | 10.715 | 10.6 | 10.45 | 10.55 | 10.595 | 10 | 10.3 | 9.615 | 10.0195 | 9.93
| Std Dev. |  | 0.24906 | 0.010488 | 0.10954 | 0.1355 | 0.05244 | 0.2645 | 0.16355 | 0.05164 | 0.10053 | 0.04 | 0.10328 | 0.054772 | 0.10488 | 0.058652 | 0.070356 | 0.040825 | 0.070899 | 0.020292 | 0.056095
| Rel.Std.Dev. |  | 2.467% | 0.1035% | 1.014% | 1.37% | 0.503% | 2.624% | 1.569% | 0.4796% | 0.9802% | 0.3738% | 0.9774% | 0.5241% | 0.9941% | 0.5538% | 0.7025% | 0.397% | 0.738% | 0.2025% | 0.5647%
| PDM3 |  | -1.936% | -1.547% | 4.913% | -3.927% | 1.27% | -2.081% | 1.27% | 4.589% | -0.3653% | 3.941% | 2.646% | 1.513% | 2.484% | 2.873% | -2.713% | -0.1062% | -6.679% | -2.638% | -3.506%
| Z-Score (Absolute) |  | 0.48 | 0.38 | 1.31 | 0.99 | 0.36 | 0.51 | 0.36 | 1.22 | 0.07 | 1.05 | 0.72 | 0.42 | 0.67 | 0.78 | 0.68 | 0 | 1.71 | 0.66 | 0.89
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 605b (Certified Value 1015 ppm) | [23 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | Various [a] | 4A*OES/AAS | 4A*OES/AAS | 4A*AAS | 4A*MS | 4A*OES | 4A*AAS | 4A*MS | 4A*MS | 4A*AAS | 4A*AAS | 4A*AAS | 5A*OES | 4A*OES | 4A*OES/MS | 4A*AAS | 0
|  |  | 07: see footer | 0
| 1 | 1 | 972.7119 | 1058 | 1036 | 1000 | 1030 | 1000 | 1020 | 1030 | 1010 | 1090 | 947.58 | 1016.4982 | 999 | 1090 | 1050 | 1038 | 990 | N.A. | 1034.114 | 988.8474 | > 100 | 998 | 1
| 1 | 2 | 968.0831 | 1019 | 1015 | 1010 | 1010 | 996 | > 100 | 1030 | 1015 | 1060 | 943.58 | 1010.3808 | 985 | 1106 | 1070 | 1039 | 999 | N.A. | 1021.912 | 1006.0468 | > 100 | 1018 | 25
| 1 | 3 | 976.6127 | 1039 | 1019 | 1010 | 1010 | 1000 | > 100 | 1005 | 1005 | 1020 | 961.81 | 1000.254 | 990 | 1112 | 1040 | 1056 | 978 | N.A. | 1023.819 | 1012.3807 | > 100 | 1018 | 16
| 1 | 4 | 983.2674 | 1017 | 1031 | 1010 | 1010 | 996 | 1020 | 1010 | 997 | 1100 | 956.99 | 997.8972 | 981 | 1124 | 1040 | 1052 | 976 | N.A. | 1023.936 | 1013.0171 | > 100 | 998 | 1015.032
| 1 | 5 | 974.7377 | 1040 | 1034 | 1010 | 1020 | 997 | 1020 | 992 | 997 | 1060 | 929.51 | 1000.254 | 991 | 1080 | 1030 | 1042 | 980 | N.A. | 1043.149 | 1022.4925 | > 100 | 1010 | 14
| 1 | 6 | 969.969 | 992 | 1035 | 1010 | 1030 | 988 | > 100 | 1050 | 991 | 1070 | 964.2 | 1013.1474 | 986 | 1110 | 1030 | 1047 | 970 | N.A. | 1027.604 | 1015.3258 | > 100 | 1022
| Mean |  | 974.2303 | 1027.5 | 1028.3333 | 1008.3333 | 1018.3333 | 996.1667 | 1020 | 1019.5 | 1002.5 | 1066.6667 | 950.6117 | 1006.4053 | 988.6667 | 1103.6667 | 1043.3333 | 1045.6667 | 982.1667 |  | 1029.089 | 1009.685 |  | 1010.6667
| Median |  | 973.7248 | 1029 | 1032.5 | 1010 | 1015 | 996.5 | 1020 | 1020 | 1001 | 1065 | 952.285 | 1005.3174 | 988 | 1108 | 1040 | 1044.5 | 979 |  | 1025.77 | 1012.6989 |  | 1014
| Std Dev. |  | 5.401 | 23.0716 | 9.0259 | 4.0825 | 9.8319 | 4.4008 | 0 | 21.0119 | 9.0719 | 28.0476 | 13.074 | 7.8891 | 6.2183 | 15.9708 | 15.0555 | 7.2847 | 10.5151 |  | 8.136 | 11.5036 |  | 10.5578
| Rel.Std.Dev. |  | 0.5544% | 2.245% | 0.8777% | 0.4049% | 0.9655% | 0.4418% | 0% | 2.061% | 0.9049% | 2.629% | 1.375% | 0.7839% | 0.629% | 1.447% | 1.443% | 0.6967% | 1.071% |  | 0.7906% | 1.139% |  | 1.045%
| PDM3 |  | -4.02% | 1.228% | 1.31% | -0.6599% | 0.3252% | -1.859% | 0.4894% | 0.4402% | -1.235% | 5.087% | -6.347% | -0.8499% | -2.597% | 8.732% | 2.788% | 3.018% | -3.238% |  | 1.385% | -0.5268% |  | -0.4301%
| Z-Score (Absolute) |  | 1.91 | 0.62 | 0.66 | 0.29 | 0.18 | 0.87 | 0.26 | 0.24 | 0.57 | 2.48 | 3.03 | 0.38 | 1.23 | 4.23 | 1.37 | 1.48 | 1.53 | NA | 0.69 | 0.23 | NA | 0.18
[a]: Lab 07: Rnd 01 - 4A*OES, Rnd 01 - 4A*OES/MS
| Analytical results for Al in OREAS 605b (Certified Value 6.26 wt.%) | [23 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 5.9405 | 6.3189 | 5.87 | 2.76 | 6.1 | 5.94 | 6.31 | 6.6 | 6.37 | 6.51 | 6.336 | 6.0367 | 5.94 | 6.14 | 6.21 | 6.8 | 6.29 | 5.95 | 6.4527 | 6.3261 | 6.29 | 5.146 | 1
| 1 | 2 | 5.9797 | 6.3192 | 6.13 | 2.76 | 6.16 | 5.93 | 6.29 | 6.6 | 6.28 | 6.69 | 6.2302 | 5.9411 | 5.9148 | 6.14 | 6.17 | 6.52 | 6.32 | 6.1 | 6.5837 | 6.2268 | 6.315 | 4.981 | #N/A
| 1 | 3 | 5.9705 | 6.2518 | 6.18 | 2.76 | 6.07 | 6.04 | 6.21 | 5.91 | 6.39 | 6.7 | 6.2682 | 5.9315 | 5.8068 | 6.13 | 6.21 | 6.63 | 6.36 | 6.16 | 6.609 | 6.2922 | 6.342 | 5.18 | 16
| 1 | 4 | 5.9441 | 6.247 | 6.16 | 2.77 | 6.31 | 6.04 | 6.23 | 6.48 | 6.26 | 6.54 | 6.4167 | 6.1005 | 5.9733 | 6.26 | 6.2 | 6.79 | 6.21 | 6.04 | 6.3184 | 6.403 | 6.308 | 5.171 | 6.2562
| 1 | 5 | 5.9412 | 5.9563 | 6.24 | 2.75 | 6.36 | 6 | 6.24 | 6.5 | 6.23 | 6.47 | 6.3542 | 6.1001 | 6.0273 | 5.91 | 6.22 | 6.85 | 6.17 | 5.94 | 6.5838 | 6.3295 | 6.369 | 5.071 | 15
| 1 | 6 | 5.9511 | 6.3313 | 5.88 | 2.75 | 6.52 | 6.52 | 6.35 | 6.48 | 6.13 | 6.67 | 6.3922 | 6.1804 | 5.8349 | 5.95 | 6.22 | 6.78 | 6.23 | 6.11 | 6.6142 | 6.2909 | 6.329 | 4.95
| Mean |  | 5.9545 | 6.2374 | 6.0767 | 2.7583 | 6.2533 | 6.0783 | 6.2717 | 6.4283 | 6.2767 | 6.5967 | 6.3329 | 6.0484 | 5.9162 | 6.0883 | 6.205 | 6.7283 | 6.2633 | 6.05 | 6.527 | 6.3114 | 6.3255 | 5.0832
| Median |  | 5.9476 | 6.2854 | 6.145 | 2.76 | 6.235 | 6.02 | 6.265 | 6.49 | 6.27 | 6.605 | 6.3451 | 6.0684 | 5.9274 | 6.135 | 6.21 | 6.785 | 6.26 | 6.07 | 6.5837 | 6.3092 | 6.322 | 5.1085
| Std Dev. |  | 0.016636 | 0.14246 | 0.16033 | 0.0075277 | 0.17409 | 0.22149 | 0.053821 | 0.26003 | 0.095429 | 0.10152 | 0.071743 | 0.098096 | 0.083367 | 0.13227 | 0.018708 | 0.12608 | 0.072019 | 0.089889 | 0.11817 | 0.05812 | 0.027761 | 0.099357
| Rel.Std.Dev. |  | 0.2794% | 2.284% | 2.639% | 0.2729% | 2.784% | 3.644% | 0.8582% | 4.045% | 1.52% | 1.539% | 1.133% | 1.622% | 1.409% | 2.173% | 0.3015% | 1.874% | 1.15% | 1.486% | 1.81% | 0.9209% | 0.4389% | 1.955%
| PDM3 |  | -4.823% | -0.3009% | -2.87% | -55.91% | -0.04654% | -2.844% | 0.2465% | 2.751% | 0.3264% | 5.441% | 1.226% | -3.322% | -5.436% | -2.684% | -0.8191% | 7.546% | 0.1133% | -3.297% | 4.327% | 0.8818% | 1.107% | -18.75%
| Z-Score (Absolute) |  | 1.17 | 0.03 | 0.68 | 14.01 | 0.03 | 0.67 | 0.11 | 0.73 | 0.13 | 1.41 | 0.35 | 0.79 | 1.32 | 0.63 | 0.16 | 1.94 | 0.07 | 0.78 | 1.13 | 0.27 | 0.32 | 4.67
| Analytical results for As in OREAS 605b (Certified Value 2454 ppm) | [23 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 2403 | 1001.9 | 2343 | 2410 | 2440 | 2510 | 2630 | 2570 | 2730 | 2503.5 | 2446.2173 | 2144.1 | 2600 | 2430 | 2511 | 2240 | 2389 | 1223.359 | 2426.9556 | 2128 | 2335.907 |  | 1
| 1 | 2 | 2359 | 950.3 | 2337 | 2440 | 2430 | 2490 | 2650 | 2540 | 2670 | 2500.4 | 2415.1211 | 2340.7 | 2600 | 2400 | 2433 | 2270 | 2273 | 1200.964 | 2462.8541 | 2094 | 2234.523 |  | 26
| 1 | 3 | 2359 | 829.5 | 2352 | 2420 | 2400 | 2470 | 2510 | 2600 | 2720 | 2480.9 | 2391.4126 | 2363.6 | 2620 | 2430 | 2465 | 2160 | 2371 | 1549.811 | 2512.0692 | 2035 | 2375.968 |  | 16
| 1 | 4 | 2300 | 811.6 | 2378 | 2460 | 2380 | 2470 | 2540 | 2480 | 2700 | 2566.7 | 2399.4053 | 2328.9 | 2620 | 2420 | 2402 | 2230 | 2295 | 1559.979 | 2506.726 | 2112 | 2286.534 |  | 2453.897
| 1 | 5 | 2190 | 842 | 2388 | 2680 | 2420 | 2470 | 2610 | 2540 | 2680 | 2560.8 | 2511.1334 | 2380.1 | 2570 | 2500 | 2461 | 2320 | 2379 | 1457.736 | 2524.7373 | 2068 | 2309.026 |  | 16
| 1 | 6 | 2385 | 1031.4 | 2387 | 2720 | 2540 | 2510 | 2590 | 2500 | 2620 | 2560.4 | 2525.3502 | 2341 | 2580 | 2440 | 2487 | 2220 | 2287 | 1232.31 | 2533.4644 | 2088 | 2376.407
| Mean |  | 2332.6667 | 911.1167 | 2364.1667 | 2521.6667 | 2435 | 2486.6667 | 2588.3333 | 2538.3333 | 2686.6667 | 2528.7833 | 2448.1066 | 2316.4 | 2598.3333 | 2436.6667 | 2459.8333 | 2240 | 2332.3333 | 1370.6932 | 2494.4678 | 2087.5 | 2319.7275
| Median |  | 2359 | 896.15 | 2365 | 2450 | 2425 | 2480 | 2600 | 2540 | 2690 | 2531.95 | 2430.6692 | 2340.85 | 2600 | 2430 | 2463 | 2235 | 2333 | 1345.023 | 2509.3976 | 2091 | 2322.4665
| Std Dev. |  | 78.094 | 95.4851 | 22.8685 | 139.7736 | 55.7674 | 19.6638 | 53.8207 | 44.0076 | 39.833 | 37.9474 | 57.6482 | 86.3999 | 20.4124 | 33.8625 | 38.5975 | 53.2917 | 52.6371 | 170.3843 | 41.1203 | 32.9226 | 54.979
| Rel.Std.Dev. |  | 3.348% | 10.48% | 0.9673% | 5.543% | 2.29% | 0.7908% | 2.079% | 1.734% | 1.483% | 1.501% | 2.355% | 3.73% | 0.7856% | 1.39% | 1.569% | 2.379% | 2.257% | 12.43% | 1.648% | 1.577% | 2.37%
| PDM3 |  | -4.94% | -62.87% | -3.657% | 2.762% | -0.7701% | 1.335% | 5.478% | 3.441% | 9.486% | 3.052% | -0.236% | -5.603% | 5.886% | -0.7022% | 0.2419% | -8.717% | -4.954% | -44.14% | 1.653% | -14.93% | -5.468%
| Z-Score (Absolute) |  | 0.67 | 9.89 | 0.47 | 0.55 | 0.01 | 0.32 | 0.98 | 0.66 | 1.62 | 0.6 | 0.07 | 0.78 | 1.05 | 0 | 0.15 | 1.28 | 0.68 | 6.91 | 0.37 | 2.26 | 0.76
| Analytical results for Au in OREAS 605b (Indicative Value 1.64 ppm) | [23 empty] | Indicative
Au | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 23 |  | ppm
|  |  | 02: 5A*MS |  | 2
2
| 1 | 1 | 02: 1.6243 |  | 1
| 1 | 2 | 02: 1.6071 |  | 5
| 1 | 3 | 02: 1.7858 |  | 16
| 1 | 4 | 02: 1.7067 |  | 1.6367
| 1 | 5 | 02: 1.681 |  | 11
| 1 | 6 | 02: 1.4156
| Mean |  | 02: 1.6367
| Median |  | 02: 1.6526
| Std Dev. |  | 02: 0.1257
| Rel.Std.Dev. |  | 02: 7.68%
| PDM3 |  | 02: 1.998e-13%
Z-Score (Absolute) | bar: 02
| Analytical results for Ba in OREAS 605b (Indicative Value 413 ppm) | [23 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  |  | 0
0
| 1 | 1 | 4401 | 21 | 223 | 1430 | 200 | 130 | 60 | 80 | 1260 | > 5000 | 54 | 1119 | 50.2 | 587 | 1849.67 | 156.7111 | 2414 | 613.396 |  |  |  |  | 1
| 1 | 2 | > 5000 | 72 | 229 | 980 | 140 | 170 | 80 | 160 | 1250 | > 5000 | 62 | 1147 | 51.6 | 557 | 1974.336 | 156.3511 | 2477 | 591.078 |  |  |  |  | 6
| 1 | 3 | 4450 | 42 | 217 | 1010 | 220 | 210 | 20 | 70 | 1340 | 4860.5143 | 62 | 1228 | 49.7 | 504 | 2064.97 | 166.8044 | 2470 | 544.614 |  |  |  |  | 16
| 1 | 4 | 4993 | 126 | 218 | 700 | 120 | 220 | 150 | 70 | 1280 | > 5000 | 107 | 1148 | 42.1 | 542 | 1668.86 | 159.5181 | 2479 | 593.077 |  |  |  |  | 413.0619
| 1 | 5 | > 5000 | 37 | 216 | 1490 | 130 | 170 | 60 | 70 | 1340 | > 5000 | 67 | 1214 | 50.3 | 542 | 2054.046 | 158.5052 | 2479 | 544.743 |  |  |  |  | 12
| 1 | 6 | > 5000 | 42 | 210 | 1480 | 1780 | 110 | 130 | 60 | 1310 | > 5000 | 190 | 1154 | 48.7 | 508 | 2430.391 | 160.4927 | 2512 | 537.106
| Mean |  | 4614.6667 | 56.6667 | 218.8333 | 1181.6667 | 431.6667 | 168.3333 | 83.3333 | 85 | 1296.6667 | 4860.5143 | 90.3333 | 1168.3333 | 48.7667 | 540 | 2007.0455 | 159.7304 | 2471.8333 | 570.669
| Median |  | 4450 | 42 | 217.5 | 1220 | 170 | 170 | 70 | 70 | 1295 | 4860.5143 | 64.5 | 1151 | 49.95 | 542 | 2014.191 | 159.0117 | 2478 | 567.9105
| Std Dev. |  | 328.561 | 37.7659 | 6.4936 | 331.0237 | 661.7678 | 43.0891 | 48.4424 | 37.2827 | 39.3277 |  | 52.3093 | 42.7863 | 3.3986 | 31.0677 | 254.9479 | 3.8132 | 31.9087 | 32.3164
| Rel.Std.Dev. |  | 7.12% | 66.65% | 2.967% | 28.01% | 153.3% | 25.6% | 58.13% | 43.86% | 3.033% |  | 57.91% | 3.662% | 6.969% | 5.753% | 12.7% | 2.387% | 1.291% | 5.663%
| PDM3 |  | 1017% | -86.28% | -47.02% | 186.1% | 4.504% | -59.25% | -79.83% | -79.42% | 213.9% | 1077% | -78.13% | 182.8% | -88.19% | 30.73% | 385.9% | -61.33% | 498.4% | 38.16%
| Z-Score (Absolute) |  | 6.7 | 0.7 | 0.43 | 1.13 | 0.09 | 0.51 | 0.65 | 0.65 | 1.31 | 7.09 | 0.64 | 1.11 | 0.71 | 0.09 | 2.47 | 0.53 | 3.22 | 0.14
| Analytical results for Be in OREAS 605b (Certified Value 0.71 ppm) | [23 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES |  | 2
2
| 1 | 1 | 0.76136 | 0.8 | 2 | < 5 | 0.79 | 0.87 | 0.79 | 0.82 | 0.78 | 0.8 | 0.69 | 0.62239 | < 1 | 0.93 | 0.9 | 0.9 | 0.61 | 0.6 | 0.607 | < 0.5 | 0.632 |  | 1
| 1 | 2 | 0.68158 | 0.8 | < 1 | < 5 | 0.79 | 0.65 | 0.7 | 0.84 | 0.76 | 0.7 | 0.73 | 0.65999 | < 1 | 0.92 | 0.9 | 0.9 | 0.74 | 0.6 | 0.604 | < 0.5 | 0.63 |  | 28
| 1 | 3 | 0.68627 | 0.8 | < 1 | < 5 | 0.87 | 0.63 | 0.68 | 0.81 | 0.73 | 0.7 | 0.79 | 0.62688 | < 1 | 0.94 | 0.8 | 0.9 | 0.68 | 0.7 | 0.623 | < 0.5 | 0.676 |  | 16
| 1 | 4 | 0.64513 | 0.8 | < 1 | < 5 | 0.74 | 0.81 | 0.85 | 0.65 | 0.7 | 0.7 | 0.75 | 0.62158 | < 1 | 0.92 | 0.8 | 0.8 | 0.73 | 0.6 | 0.587 | < 0.5 | 0.654 |  | 0.71014
| 1 | 5 | 0.72754 | 0.8 | < 1 | < 5 | 0.74 | 0.86 | 0.69 | 0.71 | 0.66 | 0.7 | 0.79 | 0.64735 | < 1 | 0.87 | 0.8 | 0.9 | 0.79 | 0.7 | 0.615 | < 0.5 | 0.643 |  | 17
| 1 | 6 | 0.69984 | 0.8 | < 1 | < 5 | 0.84 | 0.79 | 0.68 | 0.55 | 0.68 | 0.7 | 0.75 | 0.63032 | < 1 | 0.93 | 0.8 | 0.9 | 0.95 | 0.6 | 0.61 | < 0.5 | 0.621
| Mean |  | 0.70029 | 0.8 | 2 |  | 0.795 | 0.76833 | 0.73167 | 0.73 | 0.71833 | 0.71667 | 0.75 | 0.63475 |  | 0.91833 | 0.83333 | 0.88333 | 0.75 | 0.63333 | 0.60767 |  | 0.64267
| Median |  | 0.69305 | 0.8 | 2 |  | 0.79 | 0.8 | 0.695 | 0.76 | 0.715 | 0.7 | 0.75 | 0.6286 |  | 0.925 | 0.8 | 0.9 | 0.735 | 0.6 | 0.6085 |  | 0.6375
| Std Dev. |  | 0.040152 | 1.2162e-16 |  |  | 0.05244 | 0.104 | 0.071391 | 0.11472 | 0.046655 | 0.040825 | 0.037947 | 0.015517 |  | 0.024833 | 0.05164 | 0.040825 | 0.11541 | 0.05164 | 0.012127 |  | 0.019916
| Rel.Std.Dev. |  | 5.734% | 1.52e-14% |  |  | 6.596% | 13.54% | 9.757% | 15.71% | 6.495% | 5.696% | 5.06% | 2.445% |  | 2.704% | 6.197% | 4.622% | 15.39% | 8.154% | 1.996% |  | 3.099%
| PDM3 |  | -1.388% | 12.65% | 181.6% |  | 11.95% | 8.194% | 3.031% | 2.796% | 1.153% | 0.9188% | 5.613% | -10.62% |  | 29.32% | 17.35% | 24.39% | 5.613% | -10.82% | -14.43% |  | -9.502%
| Z-Score (Absolute) |  | 0.52 | NA | 0.35 | 31.14 | 1.14 | 0.67 | 0.03 | 0 | 0.21 | 0.23 | 0.35 | 1.68 | 4.05 | 3.31 | NA | NA | 0.35 | NA | 2.15 | 8.44 | 1.54
Results from laboratories 2, 16, 20 & 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 605b (Certified Value 49.6 ppm) | [23 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 50.4186 | 46 | 57.51 | 32 | 49.1 | 47.8 | 47.5 | 53.8 | 54.2 | 51.7 | 50.89 | 48.1072 | 44.55 | 39.7 | 56.7 | 52.2 | 45.1 | 54 | 50.32 | 49.3248 | < 5 | 47.792 | 1
| 1 | 2 | 50.3665 | 45.8 | 55.12 | 32 | 49 | 49.5 | 48.4 | 53.4 | 52.3 | 51.2 | 51.72 | 48.0233 | 43.76 | 46.4 | 49.3 | 50.2 | 46.8 | 56 | 50.7747 | 48.9242 | < 5 | 44.761 | 29
| 1 | 3 | 50.7205 | 46.99 | 55.04 | 33 | 48.9 | 49.4 | 48.5 | 49 | 53.1 | 51.3 | 41.96 | 47.7093 | 43.21 | 39.8 | 56.4 | 50.2 | 45.1 | 57 | 52.2588 | 59.0699 | < 5 | 47.156 | 16
| 1 | 4 | 50.5204 | 47.65 | 56.37 | 33 | 52.9 | 45.4 | 46.7 | 52.4 | 49.3 | 51.9 | 37.93 | 48.0833 | 43.96 | 39.3 | 54 | 50.5 | 47.2 | 44 | 50.5329 | 48.1535 | < 5 | 49.632 | 49.5801
| 1 | 5 | 50.2252 | 45.16 | 56.45 | 33 | 58.1 | 49.7 | 47.1 | 52.1 | 52 | 51.8 | 39.1 | 48.2018 | 44.22 | 43 | 57.5 | 51.9 | 50 | 51 | 52.8152 | 49.2946 | < 5 | 49.186 | 18
| 1 | 6 | 50.261 | 44.98 | 54.22 | 33 | 58.6 | 54.1 | 49.5 | 52.9 | 49.7 | 51.7 | 41.23 | 47.723 | 42.98 | 45.7 | 60.2 | 52.3 | 47.6 | 58 | 51.2669 | 49.0424 | < 5 | 47.25
| Mean |  | 50.4187 | 46.0967 | 55.785 | 32.6667 | 52.7667 | 49.3167 | 47.95 | 52.2667 | 51.7667 | 51.6 | 43.805 | 47.9746 | 43.78 | 42.3167 | 55.6833 | 51.2167 | 46.9667 | 53.3333 | 51.3281 | 50.6349 |  | 47.6295
| Median |  | 50.3926 | 45.9 | 55.745 | 33 | 51 | 49.45 | 47.95 | 52.65 | 52.15 | 51.7 | 41.595 | 48.0533 | 43.86 | 41.4 | 56.55 | 51.2 | 47 | 55 | 51.0208 | 49.1685 |  | 47.521
| Std Dev. |  | 0.18246 | 1.0424 | 1.2005 | 0.5164 | 4.5842 | 2.8534 | 1.0387 | 1.7178 | 1.918 | 0.28284 | 5.9922 | 0.20838 | 0.59736 | 3.1896 | 3.7102 | 1.0187 | 1.825 | 5.2026 | 1.0036 | 4.1541 |  | 1.7353
| Rel.Std.Dev. |  | 0.3619% | 2.261% | 2.152% | 1.581% | 8.688% | 5.786% | 2.166% | 3.287% | 3.705% | 0.5481% | 13.68% | 0.4344% | 1.364% | 7.537% | 6.663% | 1.989% | 3.886% | 9.755% | 1.955% | 8.204% |  | 3.643%
| PDM3 |  | 1.691% | -7.026% | 12.51% | -34.11% | 6.427% | -0.5313% | -3.288% | 5.419% | 4.41% | 4.074% | -11.65% | -3.238% | -11.7% | -14.65% | 12.31% | 3.301% | -5.271% | 7.57% | 3.526% | 2.127% |  | -3.934%
| Z-Score (Absolute) |  | 0.14 | 0.96 | 1.51 | 4.38 | 0.74 | 0.14 | 0.49 | 0.61 | 0.48 | 0.44 | 1.54 | 0.48 | 1.55 | 1.92 | 1.48 | 0.34 | 0.74 | 0.88 | 0.37 | 0.2 | 12.06 | 0.57
| Analytical results for Ca in OREAS 605b (Certified Value 0.423 wt.%) | [23 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.37474 | 0.4304 | 0.41 | 0.21 | 0.45 | 0.42 | 0.43 | 0.44 | 0.43 | 0.413 | 0.4225 | 0.42241 | 0.3959 | 0.42 | 0.44 | 0.45 | 0.41 | 0.39 | 0.42516 | 0.4242 | 0.4488 | 0.425 | 1
| 1 | 2 | 0.37272 | 0.4252 | 0.41 | 0.2 | 0.45 | 0.42 | 0.43 | 0.44 | 0.43 | 0.408 | 0.424 | 0.40703 | 0.3995 | 0.42 | 0.44 | 0.45 | 0.41 | 0.4 | 0.41855 | 0.41264 | 0.4519 | 0.408 | #N/A
| 1 | 3 | 0.37548 | 0.4218 | 0.39 | 0.2 | 0.45 | 0.43 | 0.42 | 0.43 | 0.43 | 0.409 | 0.4161 | 0.41041 | 0.392 | 0.42 | 0.43 | 0.45 | 0.414 | 0.41 | 0.42437 | 0.42308 | 0.4407 | 0.439 | 16
| 1 | 4 | 0.37522 | 0.4242 | 0.42 | 0.2 | 0.45 | 0.42 | 0.42 | 0.47 | 0.42 | 0.413 | 0.4277 | 0.41776 | 0.4089 | 0.43 | 0.44 | 0.45 | 0.403 | 0.4 | 0.40244 | 0.42698 | 0.4552 | 0.419 | 0.4235
| 1 | 5 | 0.37901 | 0.4071 | 0.42 | 0.2 | 0.44 | 0.42 | 0.43 | 0.47 | 0.43 | 0.411 | 0.41 | 0.4207 | 0.4108 | 0.4 | 0.44 | 0.46 | 0.411 | 0.39 | 0.42298 | 0.41015 | 0.45 | 0.407 | 19
| 1 | 6 | 0.37355 | 0.4273 | 0.38 | 0.21 | 0.45 | 0.45 | 0.44 | 0.47 | 0.42 | 0.405 | 0.4328 | 0.41474 | 0.391 | 0.4 | 0.44 | 0.43 | 0.407 | 0.4 | 0.43065 | 0.41094 | 0.455 | 0.371
| Mean |  | 0.37512 | 0.42267 | 0.405 | 0.20333 | 0.44833 | 0.42667 | 0.42833 | 0.45333 | 0.42667 | 0.40983 | 0.42218 | 0.41551 | 0.39968 | 0.415 | 0.43833 | 0.44833 | 0.40917 | 0.39833 | 0.42069 | 0.418 | 0.45027 | 0.4115
| Median |  | 0.37498 | 0.4247 | 0.41 | 0.2 | 0.45 | 0.42 | 0.43 | 0.455 | 0.43 | 0.41 | 0.42325 | 0.41625 | 0.3977 | 0.42 | 0.44 | 0.45 | 0.41 | 0.4 | 0.42368 | 0.41786 | 0.45095 | 0.4135
| Std Dev. |  | 0.0021715 | 0.0081623 | 0.016432 | 0.005164 | 0.0040825 | 0.012111 | 0.0075277 | 0.018619 | 0.005164 | 0.0031252 | 0.0081485 | 0.0059681 | 0.0084535 | 0.012247 | 0.0040825 | 0.0098319 | 0.0037639 | 0.0075277 | 0.0097517 | 0.0075531 | 0.0053508 | 0.023098
| Rel.Std.Dev. |  | 0.5789% | 1.931% | 4.057% | 2.54% | 0.9106% | 2.838% | 1.757% | 4.107% | 1.21% | 0.7625% | 1.93% | 1.436% | 2.115% | 2.951% | 0.9314% | 2.193% | 0.9199% | 1.89% | 2.318% | 1.807% | 1.188% | 5.613%
| PDM3 |  | -11.42% | -0.196% | -4.368% | -51.99% | 5.865% | 0.7485% | 1.142% | 7.045% | 0.7485% | -3.226% | -0.3101% | -1.886% | -5.623% | -2.006% | 3.503% | 5.865% | -3.384% | -5.942% | -0.6619% | -1.298% | 6.321% | -2.833%
| Z-Score (Absolute) |  | 3.03 | 0.23 | 0.98 | 14.79 | 1.98 | 0.5 | 0.62 | 2.33 | 0.5 | 0.65 | 0.19 | 0.26 | 1.35 | 0.3 | 1.3 | 1.98 | 0.7 | 1.44 | 0.09 | 0.09 | 2.12 | 0.54
| Analytical results for Cd in OREAS 605b (Certified Value 14 ppm) | [23 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 13.0154 | 12.2 | 15.72 | 23 | 13.8 | 12.95 | 13.45 | 14.2 | 13.3 | 13.5 | 13.91 | 13.3732 | 12.62 | 14.4 | 16 | 15.9 | 13.2 | 5 | 15.6073 | 14.0313 | 35.36 | 14.692 | 1
| 1 | 2 | 13.2289 | 11.96 | 15.18 | 23 | 14.1 | 12.6 | 13.2 | 14.9 | 13.2 | 13.3 | 13.93 | 13.6292 | 10.76 | 14.6 | 15 | 15.1 | 13.2 | 5 | 14.8798 | 14.0996 | 36.43 | 14.272 | 30
| 1 | 3 | 13.083 | 12.34 | 15.22 | 23 | 13.9 | 12.85 | 13.65 | 14 | 13.05 | 13.6 | 13.75 | 13.3361 | 12.63 | 14.4 | 16 | 15.2 | 13.5 | 5 | 15.1145 | 14.2269 | 35.62 | 15.026 | 16
| 1 | 4 | 13.4347 | 12.14 | 15.26 | 23 | 13.65 | 14.55 | 13.9 | 15.2 | 12.15 | 13.5 | 14.14 | 13.2878 | 11.3 | 14.1 | 15.5 | 15.2 | 13.2 | 5 | 14.6034 | 13.9058 | 34.97 | 14.553 | 13.965
| 1 | 5 | 13.5937 | 11.88 | 15.05 | 23 | 14.75 | 13.75 | 14.6 | 15.25 | 13.05 | 13.8 | 14.23 | 13.5397 | 13.49 | 14 | 16.5 | 15.2 | 13.4 | 5 | 14.9574 | 14.1598 | 34.21 | 14.632 | 20
| 1 | 6 | 12.966 | 12.93 | 15.14 | 23 | 14.6 | 13.7 | 14.5 | 14.7 | 12.25 | 13.9 | 14.32 | 13.1927 | 11.91 | 14.5 | 17 | 15.4 | 13.1 | 5 | 15.457 | 14.0225 | 35.2 | 15.188
| Mean |  | 13.2203 | 12.2417 | 15.2617 | 23 | 14.1333 | 13.4 | 13.8833 | 14.7083 | 12.8333 | 13.6 | 14.0467 | 13.3931 | 12.1183 | 14.3333 | 16 | 15.3333 | 13.2667 | 5 | 15.1032 | 14.0743 | 35.2983 | 14.7272
| Median |  | 13.156 | 12.17 | 15.2 | 23 | 14 | 13.325 | 13.775 | 14.8 | 13.05 | 13.55 | 14.035 | 13.3546 | 12.265 | 14.4 | 16 | 15.2 | 13.2 | 5 | 15.036 | 14.0654 | 35.28 | 14.662
| Std Dev. |  | 0.24939 | 0.37568 | 0.23583 | 0 | 0.44684 | 0.73212 | 0.56627 | 0.51616 | 0.50067 | 0.21909 | 0.21787 | 0.16257 | 0.9953 | 0.23381 | 0.70711 | 0.29439 | 0.15055 | 0 | 0.37427 | 0.11327 | 0.73347 | 0.3315
| Rel.Std.Dev. |  | 1.886% | 3.069% | 1.545% | 0% | 3.162% | 5.464% | 4.079% | 3.509% | 3.901% | 1.611% | 1.551% | 1.214% | 8.213% | 1.631% | 4.419% | 1.92% | 1.135% | 0% | 2.478% | 0.8048% | 2.078% | 2.251%
| PDM3 |  | -5.333% | -12.34% | 9.285% | 64.7% | 1.205% | -4.046% | -0.585% | 5.323% | -8.104% | -2.614% | 0.5846% | -4.095% | -13.22% | 2.637% | 14.57% | 9.798% | -5.001% | -64.2% | 8.15% | 0.7827% | 152.8% | 5.458%
| Z-Score (Absolute) |  | 0.78 | 1.68 | 1.11 | NA | 0.07 | 0.61 | 0.16 | 0.6 | 1.13 | 0.43 | 0.01 | 0.62 | 1.79 | 0.25 | 1.79 | 1.17 | 0.73 | NA | 0.96 | 0.01 | 19.6 | 0.62
Results from laboratories 4 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 605b (Certified Value 21.2 ppm) | [23 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 11 | 14 | 15 | 16 | 20 | 21 | 23 | 24 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*OES | 4A*OES/MS | [7 empty] | 1
1
| 1 | 1 | 26.4801 | 10.19 | 17.1 | 22.3 | 23.2 | 19.75 | 24.3 | 34.75 | 10 | 15.4 | 29.9 | 22.55 | 14.3 | 26.2239 | 21.8636 | [7 empty] | 1
| 1 | 2 | 26.637 | 8.35 | 18.6 | 23.7 | 22.2 | 20.7 | 23.3 | 34.84 | 12 | 17.7 | 25.2 | 22.61 | 13.2 | 25.931 | 22.7037 | [7 empty] | 31
| 1 | 3 | 26.4569 | 7.97 | 15.95 | 23.9 | 23.7 | 14 | 23.4 | 33.69 | 13 | 15.1 | 25.2 | 21.33 | 14.1 | 27.2192 | 21.483 | [7 empty] | 16
| 1 | 4 | 26.8198 | 8.92 | 16.3 | 22.6 | 22.9 | 23.4 | 23.1 | 34.55 | 16 | 15.4 | 28.2 | 21.37 | 14.1 | 25.9679 | 22.2003 | [7 empty] | 21.2325
| 1 | 5 | 25.9362 | 9.62 | 18.5 | 22.5 | 22.5 | 24.4 | 24.1 | 34.34 | 17 | 16.8 | 23.5 | 22.31 | 13.2 | 26.8748 | 20.6783 | [7 empty] | 21
| 1 | 6 | 26.5195 | 8.14 | 20.4 | 26.9 | 22.2 | 22.2 | 22.7 | 34.37 | 15 | 18.2 | 27.9 | 21.43 | 14.1 | 26.2822 | 21.9172
| Mean |  | 26.4749 | 8.865 | 17.8083 | 23.65 | 22.7833 | 20.7417 | 23.4833 | 34.4233 | 13.8333 | 16.4333 | 26.65 | 21.9333 | 13.8333 | 26.4165 | 21.8077
| Median |  | 26.4998 | 8.635 | 17.8 | 23.15 | 22.7 | 21.45 | 23.35 | 34.46 | 14 | 16.1 | 26.55 | 21.87 | 14.1 | 26.2531 | 21.8904
| Std Dev. |  | 0.2959 | 0.8855 | 1.6764 | 1.7248 | 0.59805 | 3.7152 | 0.608 | 0.41084 | 2.6394 | 1.3246 | 2.394 | 0.61883 | 0.49666 | 0.51893 | 0.68598
| Rel.Std.Dev. |  | 1.118% | 9.989% | 9.414% | 7.293% | 2.625% | 17.91% | 2.589% | 1.193% | 19.08% | 8.061% | 8.983% | 2.821% | 3.59% | 1.964% | 3.146%
| PDM3 |  | 24.69% | -58.25% | -16.13% | 11.39% | 7.304% | -2.312% | 10.6% | 62.13% | -34.85% | -22.6% | 25.52% | 3.301% | -34.85% | 24.42% | 2.709%
| Z-Score (Absolute) |  | 0.68 | 1.97 | 0.62 | 0.26 | 0.13 | 0.18 | 0.23 | 1.88 | 1.22 | 0.83 | 0.71 | 0 | 1.22 | 0.67 | 0.02
| Analytical results for Co in OREAS 605b (Certified Value 19.4 ppm) | [23 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 18.6106 | 19.6 | 20 | 20 | 18.4 | 18.8 | 18.8 | 19.1 | 19.7 | 20 | 19.4 | 19.3785 | 20.5 | 20.4 | 20 | 21.9 | 17.2 | 19 | 16.416 | 19.4243 | 19.21 | 19.309 | 1
| 1 | 2 | 19.0655 | 19 | 19.5 | 20 | 18.7 | 19.2 | 18.6 | 19.5 | 19.9 | 20 | 19 | 19.4345 | 22.1 | 19.8 | 20 | 21 | 17.1 | 18 | 16.589 | 19.1117 | 20.27 | 18.011 | 32
| 1 | 3 | 19.4282 | 19.2 | 20.7 | 20 | 18.5 | 19.3 | 19 | 18.6 | 19.6 | 20 | 19.4 | 19.3476 | 18.7 | 19.6 | 20 | 21.3 | 17.2 | 19 | 16.702 | 19.317 | 19.84 | 19.749 | 16
| 1 | 4 | 18.5403 | 19 | 20.7 | 20 | 19.8 | 19.7 | 19.6 | 19.8 | 18.2 | 20 | 19.6 | 19.2927 | 19.7 | 19.6 | 20 | 20.9 | 16.9 | 18 | 16.571 | 19.2257 | 19.93 | 19.686 | 19.3561
| 1 | 5 | 18.5587 | 18.1 | 20.6 | 20 | 18.6 | 19.3 | 19.1 | 19.6 | 19.3 | 20 | 19.3 | 19.2663 | 21.1 | 19.3 | 25 | 21.4 | 17.7 | 18 | 17.222 | 19.1119 | 19.96 | 19.572 | 22
| 1 | 6 | 19.1199 | 19.6 | 21.1 | 20 | 18.6 | 18.8 | 19.2 | 18.8 | 18.6 | 20 | 19.9 | 19.1949 | 18.8 | 19.7 | 20 | 21.7 | 16.9 | 18 | 16.559 | 18.7744 | 19.06 | 20.233
| Mean |  | 18.8872 | 19.0833 | 20.4333 | 20 | 18.7667 | 19.1833 | 19.05 | 19.2333 | 19.2167 | 20 | 19.4333 | 19.3191 | 20.15 | 19.7333 | 20.8333 | 21.3667 | 17.1667 | 18.3333 | 16.6765 | 19.1608 | 19.7117 | 19.4267
| Median |  | 18.838 | 19.1 | 20.65 | 20 | 18.6 | 19.25 | 19.05 | 19.3 | 19.45 | 20 | 19.4 | 19.3201 | 20.1 | 19.65 | 20 | 21.35 | 17.15 | 18 | 16.58 | 19.1688 | 19.885 | 19.629
| Std Dev. |  | 0.36971 | 0.55287 | 0.5785 | 0 | 0.5164 | 0.34303 | 0.34496 | 0.4761 | 0.67355 | 0 | 0.30111 | 0.085516 | 1.3383 | 0.36697 | 2.0412 | 0.38816 | 0.29439 | 0.5164 | 0.28237 | 0.22457 | 0.47199 | 0.75636
| Rel.Std.Dev. |  | 1.957% | 2.897% | 2.831% | 0% | 2.752% | 1.788% | 1.811% | 2.475% | 3.505% | 0% | 1.549% | 0.4427% | 6.642% | 1.86% | 9.798% | 1.817% | 1.715% | 2.817% | 1.693% | 1.172% | 2.394% | 3.893%
| PDM3 |  | -2.423% | -1.409% | 5.565% | 3.326% | -3.045% | -0.8927% | -1.582% | -0.6344% | -0.7205% | 3.326% | 0.3989% | -0.1914% | 4.101% | 1.949% | 7.632% | 10.39% | -11.31% | -5.284% | -13.84% | -1.009% | 1.837% | 0.3644%
| Z-Score (Absolute) |  | 0.83 | 0.35 | 2.98 | NA | 1.13 | 0.1 | 0.43 | 0.02 | 0.02 | NA | 0.51 | 0.23 | 2.28 | 1.26 | NA | 5.29 | 5.08 | NA | 6.29 | 0.16 | 1.2 | 0.5
Results from laboratories 4, 10, 16 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 605b (Certified Value 30.3 ppm) | [23 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
1
| 1 | 1 | 27 | 26 | 36 | 29 | 31 | 32 | 34 | 32 | 30 | 36 | 25.6545 | 25 | 21.3 | 20 | 29 | 36 | 30 | 33.359 | 31.5513 | 37.58 | 26.274 |  | 1
| 1 | 2 | 28 | 25 | 37 | 29 | 31 | 31 | 34 | 30 | 29 | 35 | 23.1982 | 27 | 21 | 20 | 28 | 36.1 | 29 | 33.759 | 31.1619 | 37.02 | 24.652 |  | 33
| 1 | 3 | 28 | 27 | 36 | 30 | 30 | 31 | 30 | 32 | 28 | 38 | 24.6298 | 27 | 20.8 | 10 | 28 | 39.6 | 31 | 33.819 | 30.4673 | 36.48 | 25.019 |  | 16
| 1 | 4 | 28 | 26 | 36 | 32 | 31 | 31 | 33 | 30 | 29 | 37 | 26.2433 | 28 | 20.8 | 10 | 30 | 36.1 | 29 | 33.631 | 34.5887 | 37.45 | 25.624 |  | 30.3356
| 1 | 5 | 27 | 27 | 37 | 30 | 31 | 31 | 33 | 31 | 28 | 35 | 23.6881 | 28 | 21.5 | 10 | 29 | 37.1 | 30 | 34.146 | 31.836 | 36.79 | 29.95 |  | 23
| 1 | 6 | 28 | 26 | 36 | 30 | 33 | 32 | 34 | 30 | 29 | 35 | 25.1518 | 25 | 20.3 | 20 | 28 | 37.2 | 29 | 33.637 | 31.8127 | 37.57 | 25.076
| Mean |  | 27.6667 | 26.1667 | 36.3333 | 30 | 31.1667 | 31.3333 | 33 | 30.8333 | 28.8333 | 36 | 24.761 | 26.6667 | 20.95 | 15 | 28.6667 | 37.0167 | 29.6667 | 33.7252 | 31.903 | 37.1483 | 26.0992
| Median |  | 28 | 26 | 36 | 30 | 31 | 31 | 33.5 | 30.5 | 29 | 35.5 | 24.8908 | 27 | 20.9 | 15 | 28.5 | 36.6 | 29.5 | 33.698 | 31.682 | 37.235 | 25.35
| Std Dev. |  | 0.5164 | 0.75277 | 0.5164 | 1.0954 | 0.98319 | 0.5164 | 1.5492 | 0.98319 | 0.75277 | 1.2649 | 1.1627 | 1.3663 | 0.42308 | 5.4772 | 0.8165 | 1.3732 | 0.8165 | 0.25991 | 1.4114 | 0.45753 | 1.9694
| Rel.Std.Dev. |  | 1.866% | 2.877% | 1.421% | 3.651% | 3.155% | 1.648% | 4.695% | 3.189% | 2.611% | 3.514% | 4.696% | 5.123% | 2.019% | 36.51% | 2.848% | 3.71% | 2.752% | 0.7707% | 4.424% | 1.232% | 7.546%
| PDM3 |  | -8.798% | -13.74% | 19.77% | -1.106% | 2.739% | 3.289% | 8.783% | 1.641% | -4.952% | 18.67% | -18.38% | -12.09% | -30.94% | -50.55% | -5.502% | 22.02% | -2.205% | 11.17% | 5.167% | 22.46% | -13.97%
| Z-Score (Absolute) |  | 0.47 | 0.78 | 1.28 | 0 | 0.24 | 0.27 | 0.61 | 0.17 | 0.24 | 1.21 | 1.06 | 0.67 | 1.83 | 3.03 | 0.27 | 1.42 | 0.07 | 0.75 | 0.38 | 1.45 | 0.79
| Analytical results for Cs in OREAS 605b (Certified Value 2.31 ppm) | [23 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | [5 empty] | 2
3
| 1 | 1 | 2.8245 | 2.4 | 2.2 | 2.27 | 2.31 | 2.39 | 2.39 | 2.38 | 2.3 | 2.46 | 2.3786 | 1.8 | 2.14 | 2.3 | 2.7 | 2.16 | 2.25 | [5 empty] | 1
| 1 | 2 | 2.7567 | 2.3 | 2.3 | 2.32 | 2.29 | 2.3 | 2.54 | 2.43 | 2.3 | 2.38 | 2.3538 | 1.8 | 2.16 | 2.1 | 2.59 | 2.07 | 2.225 | [5 empty] | 34
| 1 | 3 | 2.7654 | 2.3 | 2.3 | 2.31 | 2.38 | 2.36 | 2.4 | 2.33 | 2.4 | 2.46 | 2.31 | 1.6 | 2.12 | 2.2 | 2.56 | 2.11 | 2.279 | [5 empty] | 16
| 1 | 4 | 2.6785 | 2.3 | 2.3 | 2.33 | 2.27 | 2.43 | 2.55 | 2.25 | 2.3 | 2.47 | 2.4181 | 1.9 | 2.11 | 2.3 | 2.44 | 2.11 | 2.215 | [5 empty] | 2.3102
| 1 | 5 | 2.7365 | 2.2 | 2.3 | 2.25 | 2.35 | 2.39 | 2.54 | 2.41 | 2.4 | 2.39 | 2.463 | 1.7 | 2.05 | 2.3 | 2.59 | 2.16 | 2.344 | [5 empty] | 24
| 1 | 6 | 2.8123 | 2.4 | 2 | 2.3 | 2.38 | 2.39 | 2.48 | 2.24 | 2.4 | 2.38 | 2.43 | 1.9 | 2.14 | 2.4 | 2.64 | 2.07 | 2.226
| Mean |  | 2.7623 | 2.3167 | 2.2333 | 2.2967 | 2.33 | 2.3767 | 2.4833 | 2.34 | 2.35 | 2.4233 | 2.3923 | 1.7833 | 2.12 | 2.2667 | 2.5867 | 2.1133 | 2.2565
| Median |  | 2.7611 | 2.3 | 2.3 | 2.305 | 2.33 | 2.39 | 2.51 | 2.355 | 2.35 | 2.425 | 2.3983 | 1.8 | 2.13 | 2.3 | 2.59 | 2.11 | 2.238
| Std Dev. |  | 0.053109 | 0.075277 | 0.12111 | 0.030768 | 0.046904 | 0.043665 | 0.072847 | 0.080994 | 0.054772 | 0.044121 | 0.055728 | 0.1169 | 0.038471 | 0.10328 | 0.087101 | 0.040332 | 0.048682
| Rel.Std.Dev. |  | 1.923% | 3.249% | 5.423% | 1.34% | 2.013% | 1.837% | 2.933% | 3.461% | 2.331% | 1.821% | 2.33% | 6.555% | 1.815% | 4.556% | 3.367% | 1.908% | 2.157%
| PDM3 |  | 19.57% | 0.2785% | -3.329% | -0.5872% | 0.8557% | 2.876% | 7.493% | 1.289% | 1.721% | 4.896% | 3.55% | -22.81% | -8.234% | -1.886% | 11.97% | -8.523% | -2.326%
| Z-Score (Absolute) |  | 3.97 | 0.12 | 0.89 | 0.31 | 0 | 0.43 | 1.41 | 0.09 | 0.18 | 0.86 | 0.57 | 5.02 | 1.93 | 0.58 | 2.35 | 1.99 | 0.67
| Analytical results for Cu in OREAS 605b (Certified Value 5.03 wt.%) | [23 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | Various [a] | 4A*OES/AAS | 4A*OES/AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*AAS | 4A*MS | 4A*OES | 4A*AAS | 4A*OES | PF*OES | 5A*OES | 4A*OES | 4A*OES/MS | 4A*AAS | 2
|  |  | 07: see footer | 3
| 1 | 1 | 5.0928 | 5.2826 | 5.03 | 5.27 | 5.22 | 4.89 | 5.03 | 5.05 | 5.12 | 5.34 | 5.1551 | 4.9802 | 4.945 | 5.58 | 5.1 | 4.929 | 5.04 | 4.97 | 5.0171 | 4.8232 | > 1 | 5.08 | 1
| 1 | 2 | 5.1227 | 5.1233 | 4.906 | 5.28 | 5.23 | 4.96 | > 1 | 4.98 | 5.15 | 5.3 | 5.1636 | 5.0792 | 4.953 | 5.61 | 5.08 | 4.888 | 5.06 | 4.89 | 5.0497 | 4.8853 | > 1 | 5.05 | 35
| 1 | 3 | 5.1628 | 5.2211 | 4.917 | 5.27 | 5.34 | 4.99 | > 1 | 5.15 | 5.1 | 5.31 | 5.1596 | 4.9849 | 5.015 | 5.44 | 5.06 | 4.963 | 5.14 | 4.86 | 5.0831 | 4.909 | > 1 | 5.07 | 16
| 1 | 4 | 5.0776 | 4.7926 | 5.001 | 5.29 | 4.99 | 4.94 | 5.01 | 4.88 | 5.04 | 5.45 | 5.1476 | 5.0334 | 5.031 | 5.47 | 5.14 | 4.918 | 5.12 | 4.68 | 4.9592 | 4.8986 | > 1 | 5.03 | 5.0263
| 1 | 5 | 5.1396 | 5.1809 | 5.03 | 5.29 | 5.11 | 4.91 | 5.05 | 4.8 | 5.06 | 5.28 | 5.1663 | 5.0829 | 5.139 | 5.49 | 5.1 | 4.834 | 5.03 | 4.84 | 5.0655 | 4.9195 | > 1 | 4.95 | 25
| 1 | 6 | 5.0577 | 4.9962 | 5.02 | 5.29 | 5.14 | 4.97 | > 1 | 5.08 | 5.04 | 5.3 | 5.03 | 5.0098 | 4.901 | 5.42 | 5.06 | 4.899 | 5.09 | 4.84 | 4.8893 | 4.8943 | > 1 | 5.08
| Mean |  | 5.1089 | 5.0995 | 4.984 | 5.2817 | 5.1717 | 4.9433 | 5.03 | 4.99 | 5.085 | 5.33 | 5.137 | 5.0284 | 4.9973 | 5.5017 | 5.09 | 4.9052 | 5.08 | 4.8467 | 5.0107 | 4.8883 |  | 5.0433
| Median |  | 5.1077 | 5.1521 | 5.0105 | 5.285 | 5.18 | 4.95 | 5.03 | 5.015 | 5.08 | 5.305 | 5.1574 | 5.0216 | 4.984 | 5.48 | 5.09 | 4.9085 | 5.075 | 4.85 | 5.0334 | 4.8964 |  | 5.06
| Std Dev. |  | 0.039721 | 0.17909 | 0.057254 | 0.0098319 | 0.1199 | 0.037771 | 0.02 | 0.13084 | 0.045497 | 0.061968 | 0.05285 | 0.04502 | 0.084261 | 0.07679 | 0.030332 | 0.043531 | 0.044272 | 0.095009 | 0.073773 | 0.034053 |  | 0.049666
| Rel.Std.Dev. |  | 0.7775% | 3.512% | 1.149% | 0.1862% | 2.318% | 0.7641% | 0.3976% | 2.622% | 0.8947% | 1.163% | 1.029% | 0.8953% | 1.686% | 1.396% | 0.5959% | 0.8875% | 0.8715% | 1.96% | 1.472% | 0.6966% |  | 0.9848%
| PDM3 |  | 1.644% | 1.456% | -0.8407% | 5.082% | 2.893% | -1.65% | 0.07452% | -0.7213% | 1.169% | 6.043% | 2.204% | 0.04274% | -0.5754% | 9.459% | 1.268% | -2.409% | 1.069% | -3.573% | -0.3101% | -2.744% |  | 0.3398%
| Z-Score (Absolute) |  | 0.74 | 0.64 | 0.67 | 2.71 | 1.46 | 1.14 | 0.15 | 0.61 | 0.47 | 3.26 | 1.06 | 0.17 | 0.52 | 5.21 | 0.53 | 1.57 | 0.42 | 2.24 | 0.37 | 1.76 | NA | 0
[a]: Lab 07: Rnd 01 - 4A*OES, Rnd 01 - 4A*OES/MS
| Analytical results for Dy in OREAS 605b (Certified Value 1.34 ppm) | [23 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 23 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 5A*MS | 2
2
| 1 | 1 | 01: 1.6391; 02: 1.5; 03: 1.28; 04: 1.4; 05: 1.25; 06: 1.248 | 1
| 1 | 2 | 01: 1.5731; 02: 1.4; 03: 1.26; 04: 1.4; 05: 1.1; 06: 1.195 | 12
| 1 | 3 | 01: 1.5784; 02: 1.3; 03: 1.3; 04: 1.5; 05: 1.3; 06: 1.238 | 16
| 1 | 4 | 01: 1.5703; 02: 1.1; 03: 1.22; 04: 1.3; 05: 1.25; 06: 1.185 | 1.3442
| 1 | 5 | 01: 1.6317; 02: 1.2; 03: 1.3; 04: 1.4; 05: 1.25; 06: 1.238 | 26
| 1 | 6 | 01: 1.6411; 02: 1.2; 03: 1.28; 04: 1.5; 05: 1.3; 06: 1.195
| Mean |  | 01: 1.6056; 02: 1.2833; 03: 1.2733; 04: 1.4167; 05: 1.2417; 06: 1.2165
| Median |  | 01: 1.6051; 02: 1.25; 03: 1.28; 04: 1.4; 05: 1.25; 06: 1.2165
| Std Dev. |  | 01: 0.034952; 02: 0.1472; 03: 0.030111; 04: 0.075277; 05: 0.073598; 06: 0.027689
| Rel.Std.Dev. |  | 01: 2.177%; 02: 11.47%; 03: 2.365%; 04: 5.314%; 05: 5.927%; 06: 2.276%
| PDM3 |  | 01: 19.44%; 02: -4.531%; 03: -5.275%; 04: 5.388%; 05: -7.631%; 06: -9.503%
| Z-Score (Absolute) |  | 01: 4.48; 02: 0.07; 03: 0.07; 04: 1.89; 05: 0.5; 06: 0.85
| Analytical results for Er in OREAS 605b (Certified Value 0.64 ppm) | [23 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 23 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 5A*MS | 2
3
| 1 | 1 | 01: 0.85583; 02: 0.6; 03: 0.6; 04: 0.7; 05: 0.65; 06: 0.586 | 1
| 1 | 2 | 01: 0.84259; 02: 0.6; 03: 0.62; 04: 0.7; 05: 0.55; 06: 0.575 | 13
| 1 | 3 | 01: 0.87674; 02: 0.8; 03: 0.62; 04: 0.7; 05: 0.65; 06: 0.587 | 16
| 1 | 4 | 01: 0.85623; 02: 0.5; 03: 0.65; 04: 0.7; 05: 0.65; 06: 0.6 | 0.639
| 1 | 5 | 01: 0.91287; 02: 0.7; 03: 0.67; 04: 0.7; 05: 0.6; 06: 0.612 | 27
| 1 | 6 | 01: 0.93515; 02: 0.6; 03: 0.66; 04: 0.7; 05: 0.7; 06: 0.59
| Mean |  | 01: 0.8799; 02: 0.63333; 03: 0.63667; 04: 0.7; 05: 0.63333; 06: 0.59167
| Median |  | 01: 0.86648; 02: 0.6; 03: 0.635; 04: 0.7; 05: 0.65; 06: 0.5885
| Std Dev. |  | 01: 0.036555; 02: 0.10328; 03: 0.027325; 04: 1.2162e-16; 05: 0.05164; 06: 0.012785
| Rel.Std.Dev. |  | 01: 4.154%; 02: 16.31%; 03: 4.292%; 04: 1.737e-14%; 05: 8.154%; 06: 2.161%
| PDM3 |  | 01: 37.7%; 02: -0.8868%; 03: -0.3652%; 04: 9.546%; 05: -0.8868%; 06: -7.407%
| Z-Score (Absolute) |  | 01: 7.34; 02: 0.05; 03: 0.05; 04: 1.95; 05: 0.05; 06: 1.3
| Analytical results for Eu in OREAS 605b (Certified Value 0.46 ppm) | [23 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 23 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 5A*MS | 2
2
| 1 | 1 | 01: 0.95308; 02: 0.5; 03: 0.46; 04: 0.5; 05: 0.45; 06: 0.424 | 1
| 1 | 2 | 01: 0.99345; 02: 0.5; 03: 0.46; 04: 0.5; 05: 0.4; 06: 0.414 | 14
| 1 | 3 | 01: 0.92526; 02: 0.4; 03: 0.48; 04: 0.6; 05: 0.4; 06: 0.422 | 16
| 1 | 4 | 01: 1.0924; 02: 0.3; 03: 0.48; 04: 0.6; 05: 0.4; 06: 0.424 | 0.45667
| 1 | 5 | 01: 0.91629; 02: 0.4; 03: 0.46; 04: 0.6; 05: 0.45; 06: 0.44 | 28
| 1 | 6 | 01: 1.0178; 02: 0.3; 03: 0.47; 04: 0.6; 05: 0.45; 06: 0.416
| Mean |  | 01: 0.98305; 02: 0.4; 03: 0.46833; 04: 0.56667; 05: 0.425; 06: 0.42333
| Median |  | 01: 0.97327; 02: 0.4; 03: 0.465; 04: 0.6; 05: 0.425; 06: 0.423
| Std Dev. |  | 01: 0.066292; 02: 0.089443; 03: 0.0098319; 04: 0.05164; 05: 0.027386; 06: 0.0091797
| Rel.Std.Dev. |  | 01: 6.744%; 02: 22.36%; 03: 2.099%; 04: 9.113%; 05: 6.444%; 06: 2.168%
| PDM3 |  | 01: 115.3%; 02: -12.41%; 03: 2.555%; 04: 24.09%; 05: -6.934%; 06: -7.299%
| Z-Score (Absolute) |  | 01: 10.33; 02: 0.9; 03: 0.42; 04: 2.31; 05: 0.42; 06: 0.45
| Analytical results for Fe in OREAS 605b (Certified Value 7.14 wt.%) | [23 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 6.9351 | 7.4 | 7.14 | 7.48 | 7.01 | 6.79 | 7.01 | 6.99 | 7.1 | 7.71 | 7.15 | 7.6765 | 6.5537 | 6.93 | 7.24 | 7.16 | 6.88 | 6.73 | 7.3058 | 6.8143 | 7.229 | 7.116 | 1
| 1 | 2 | 6.9286 | 7.34 | 7.27 | 7.4 | 7.09 | 6.89 | 6.96 | 7.01 | 7.02 | 7.83 | 7.08 | 7.3622 | 7.1384 | 6.89 | 7.21 | 6.98 | 6.85 | 6.87 | 7.3912 | 6.632 | 7.21 | 6.577 | #N/A
| 1 | 3 | 6.9542 | 7.27 | 7.31 | 7.44 | 7.02 | 6.87 | 6.9 | 7 | 7.13 | 7.89 | 7.05 | 7.4891 | 7.1919 | 7.11 | 7.2 | 7 | 6.93 | 7.01 | 7.4198 | 6.8927 | 7.183 | 6.882 | 16
| 1 | 4 | 6.9539 | 7.29 | 7.4 | 7.52 | 7.14 | 6.78 | 6.93 | 7.5 | 6.86 | 7.85 | 7.31 | 7.4077 | 7.1205 | 7.12 | 7.31 | 7.11 | 6.72 | 6.79 | 7.5224 | 6.881 | 7.181 | 6.91 | 7.1376
| 1 | 5 | 6.965 | 7 | 7.4 | 7.48 | 7.29 | 6.83 | 6.93 | 7.55 | 6.95 | 7.71 | 7.2 | 7.6041 | 7.2888 | 6.74 | 7.24 | 7.1 | 6.72 | 6.66 | 7.3599 | 6.7401 | 7.208 | 7.046 | 29
| 1 | 6 | 6.9631 | 7.36 | 7.17 | 7.5 | 7.44 | 7.38 | 7.04 | 7.5 | 6.88 | 7.82 | 7.23 | 7.4872 | 7.2501 | 6.58 | 7.32 | 7.09 | 6.76 | 6.83 | 7.3496 | 6.7208 | 7.211 | 6.817
| Mean |  | 6.95 | 7.2767 | 7.2817 | 7.47 | 7.165 | 6.9233 | 6.9617 | 7.2583 | 6.99 | 7.8017 | 7.17 | 7.5045 | 7.0906 | 6.895 | 7.2533 | 7.0733 | 6.81 | 6.815 | 7.3915 | 6.7802 | 7.2037 | 6.8913
| Median |  | 6.954 | 7.315 | 7.29 | 7.48 | 7.115 | 6.85 | 6.945 | 7.255 | 6.985 | 7.825 | 7.175 | 7.4881 | 7.1651 | 6.91 | 7.24 | 7.095 | 6.805 | 6.81 | 7.3756 | 6.7772 | 7.209 | 6.896
| Std Dev. |  | 0.014896 | 0.14348 | 0.11089 | 0.043359 | 0.16885 | 0.22783 | 0.053448 | 0.28365 | 0.11243 | 0.074944 | 0.096954 | 0.11801 | 0.27071 | 0.2104 | 0.050465 | 0.069186 | 0.088994 | 0.12095 | 0.074878 | 0.10106 | 0.018414 | 0.1893
| Rel.Std.Dev. |  | 0.2143% | 1.972% | 1.523% | 0.5804% | 2.357% | 3.291% | 0.7677% | 3.908% | 1.608% | 0.9606% | 1.352% | 1.573% | 3.818% | 3.052% | 0.6957% | 0.9781% | 1.307% | 1.775% | 1.013% | 1.491% | 0.2556% | 2.747%
| PDM3 |  | -2.629% | 1.948% | 2.018% | 4.656% | 0.3834% | -3.002% | -2.465% | 1.691% | -2.068% | 9.303% | 0.4534% | 5.139% | -0.6595% | -3.399% | 1.621% | -0.9009% | -4.59% | -4.52% | 3.556% | -5.008% | 0.9251% | -3.451%
| Z-Score (Absolute) |  | 0.7 | 0.58 | 0.6 | 1.34 | 0.15 | 0.8 | 0.65 | 0.51 | 0.54 | 2.64 | 0.17 | 1.48 | 0.15 | 0.91 | 0.49 | 0.21 | 1.25 | 1.23 | 1.03 | 1.36 | 0.3 | 0.93
| Analytical results for Ga in OREAS 605b (Certified Value 32.8 ppm) | [23 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 33.5879 | 32.1 | 34.48 | 35 | 31.8 | 30.5 | 29.8 | 34.1 | 34.5 | 34.9 | 32.87 | 31.9566 | 28 | 34.1 | 32.4 | 36 | 27.1 | 27.659 | 29.4214 |  |  |  | 1
| 1 | 2 | 33.9362 | 30.9 | 33.48 | 35 | 32.5 | 31.1 | 29.9 | 34.9 | 35.1 | 32.5 | 33.17 | 31.7457 | 35 | 33.3 | 29.2 | 34.5 | 27.6 | 27.779 | 28.6939 |  |  |  | 39
| 1 | 3 | 33.8763 | 31.1 | 31.13 | 35 | 31.9 | 31.1 | 29.7 | 31.7 | 33.9 | 32.8 | 32.24 | 31.5812 | 34 | 33.8 | 32.8 | 34.4 | 27.2 | 28.252 | 28.7857 |  |  |  | 16
| 1 | 4 | 33.5774 | 30.9 | 33.46 | 36 | 34.8 | 32.6 | 30 | 35.9 | 33.2 | 32.9 | 33.85 | 31.3768 | 32 | 33.8 | 32.8 | 33.8 | 27.2 | 27.731 | 28.525 |  |  |  | 32.7744
| 1 | 5 | 33.5218 | 30.1 | 33.27 | 36 | 31.3 | 32.4 | 30.4 | 35.5 | 34.2 | 33.1 | 32.58 | 32.242 | 35 | 32.6 | 33.4 | 34.8 | 27.8 | 28.787 | 29.5363 |  |  |  | 30
| 1 | 6 | 33.988 | 32.4 | 32.96 | 36 | 31.7 | 32.7 | 30.9 | 34.7 | 32.1 | 32.1 | 34.06 | 31.4502 | 31 | 34.1 | 33.2 | 35.2 | 26.9 | 27.957 | 30.4723
| Mean |  | 33.7479 | 31.25 | 33.13 | 35.5 | 32.3333 | 31.7333 | 30.1167 | 34.4667 | 33.8333 | 33.05 | 33.1283 | 31.7254 | 32.5 | 33.6167 | 32.3 | 34.7833 | 27.3 | 28.0275 | 29.2391
| Median |  | 33.7321 | 31 | 33.365 | 35.5 | 31.85 | 31.75 | 29.95 | 34.8 | 34.05 | 32.85 | 33.02 | 31.6634 | 33 | 33.8 | 32.8 | 34.65 | 27.2 | 27.868 | 29.1036
| Std Dev. |  | 0.20756 | 0.85264 | 1.1048 | 0.54772 | 1.2691 | 0.94375 | 0.4535 | 1.4935 | 1.0577 | 0.97108 | 0.71373 | 0.32827 | 2.7386 | 0.57764 | 1.5582 | 0.75476 | 0.33466 | 0.42841 | 0.72871
| Rel.Std.Dev. |  | 0.615% | 2.728% | 3.335% | 1.543% | 3.925% | 2.974% | 1.506% | 4.333% | 3.126% | 2.938% | 2.154% | 1.035% | 8.427% | 1.718% | 4.824% | 2.17% | 1.226% | 1.529% | 2.492%
| PDM3 |  | 2.971% | -4.651% | 1.085% | 8.316% | -1.346% | -3.176% | -8.109% | 5.164% | 3.231% | 0.841% | 1.08% | -3.2% | -0.8371% | 2.57% | -1.447% | 6.13% | -16.7% | -14.48% | -10.79%
| Z-Score (Absolute) |  | 0.67 | 0.68 | 0.34 | 1.62 | 0.09 | 0.41 | 1.29 | 1.06 | 0.72 | 0.3 | 0.34 | 0.42 | 0 | 0.6 | 0.11 | 1.23 | 2.81 | 2.42 | 1.76
| Analytical results for Gd in OREAS 605b (Certified Value 1.77 ppm) | [23 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 23 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 5A*MS |  |  |  | 2
2
| 1 | 1 | 01: 2.0804; 02: 1.5; 03: 1.73; 04: 1.7; 05: 1.6; 06: 1.62 |  |  |  | 1
| 1 | 2 | 01: 2.1787; 02: 1.5; 03: 1.89; 04: 1.8; 05: 1.4; 06: 1.624 |  |  |  | 15
| 1 | 3 | 01: 2.011; 02: 2.1; 03: 1.88; 04: 1.8; 05: 1.6; 06: 1.649 |  |  |  | 16
| 1 | 4 | 01: 2.0821; 02: 2; 03: 1.88; 04: 1.8; 05: 1.4; 06: 1.597 |  |  |  | 1.7684
| 1 | 5 | 01: 2.0813; 02: 1.8; 03: 1.78; 04: 1.9; 05: 1.6; 06: 1.737 |  |  |  | 31
| 1 | 6 | 01: 2.095; 02: 1.3; 03: 1.81; 04: 2; 05: 1.6; 06: 1.646
| Mean |  | 01: 2.0881; 02: 1.7; 03: 1.8283; 04: 1.8333; 05: 1.5333; 06: 1.6455
| Median |  | 01: 2.0817; 02: 1.65; 03: 1.845; 04: 1.8; 05: 1.6; 06: 1.635
| Std Dev. |  | 01: 0.053551; 02: 0.31623; 03: 0.065549; 04: 0.10328; 05: 0.10328; 06: 0.048682
| Rel.Std.Dev. |  | 01: 2.565%; 02: 18.6%; 03: 3.585%; 04: 5.633%; 05: 6.736%; 06: 2.958%
| PDM3 |  | 01: 18.08%; 02: -3.867%; 03: 3.39%; 04: 3.673%; 05: -13.29%; 06: -6.949%
| Z-Score (Absolute) |  | 01: 2.33; 02: 0.46; 03: 0.46; 04: 0.5; 05: 1.66; 06: 0.85
| Analytical results for Ge in OREAS 605b (Certified Value 0.29 ppm) | [23 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 23 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 5A*MS | [8 empty] | 2
2
| 1 | 1 | 0.27336 | 6.4 | 0.4 | 0.34 | 0.26 | 0.36 | 0.32 | 6.8 | 7.5 | 3.8697 | < 0.05 | < 0.05 | 0.4 | 0.251 | [8 empty] | 1
| 1 | 2 | 0.2645 | 6.7 | 0.35 | 0.51 | 0.26 | 0.36 | 0.28 | 6.1 | 7.7 | 3.8393 | < 0.05 | < 0.05 | 0.5 | 0.247 | [8 empty] | 16
| 1 | 3 | 0.26655 | 6.2 | 0.37 | 0.31 | 0.27 | 0.33 | 0.26 | 6.3 | 7.2 | 3.5449 | < 0.05 | < 0.05 | 0.4 | 0.25 | [8 empty] | 16
| 1 | 4 | 0.27594 | 5.9 | 0.25 | 0.54 | 0.28 | 0.33 | 0.25 | 6.1 | 7.2 | 3.6039 | < 0.05 | < 0.05 | 0.5 | 0.286 | [8 empty] | 0.29383
| 1 | 5 | 0.26863 | 5.7 | 0.33 | 0.5 | 0.39 | 0.38 | 0.25 | 6.1 | 7.4 | 3.7608 | < 0.05 | < 0.05 | 0.4 | 0.266 | [8 empty] | 32
| 1 | 6 | 0.28399 | 6 | 0.32 | 0.83 | 0.33 | 0.28 | 0.27 | 6.1 | 7.4 | 3.6202 | < 0.05 | < 0.05 | 0.3 | 0.275
| Mean |  | 0.27216 | 6.15 | 0.33667 | 0.505 | 0.29833 | 0.34 | 0.27167 | 6.25 | 7.4 | 3.7065 |  |  | 0.41667 | 0.2625
| Median |  | 0.271 | 6.1 | 0.34 | 0.505 | 0.275 | 0.345 | 0.265 | 6.1 | 7.4 | 3.6905 |  |  | 0.4 | 0.2585
| Std Dev. |  | 0.0071889 | 0.36194 | 0.051251 | 0.18555 | 0.051929 | 0.035214 | 0.026394 | 0.28107 | 0.18974 | 0.13513 |  |  | 0.075277 | 0.015808
| Rel.Std.Dev. |  | 2.641% | 5.885% | 15.22% | 36.74% | 17.41% | 10.36% | 9.716% | 4.497% | 2.564% | 3.646% |  |  | 18.07% | 6.022%
| PDM3 |  | -7.375% | 1993% | 14.58% | 71.87% | 1.532% | 15.71% | -7.544% | 2027% | 2418% | 1161% |  |  | 41.8% | -10.66%
| Z-Score (Absolute) |  | 0.17 | NA | 0.65 | 2.79 | 0.17 | 0.7 | 0.17 | NA | NA | 43.46 | 3.31 | 3.31 | NA | 0.29
Results from laboratories 2, 10, 11 & 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 605b (Certified Value 2.27 ppm) | [23 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | [6 empty] | 2
3
| 1 | 1 | 2.3 | 2.28 | 2.2 | 1.9 | 2.2 | 2.4 | 2.4 | 2.3 | 2.43 | 2.1837 | 2.15 | 2 | 2.6 | 2.78 | 2.2 | 3.071 | [6 empty] | 1
| 1 | 2 | 2.3 | 2.14 | 2.2 | 2.2 | 2.3 | 2.5 | 2.4 | 2.3 | 2.39 | 2.078 | 2.05 | 2.3 | 2.2 | 2.74 | 2.13 | 3.088 | [6 empty] | 25
| 1 | 3 | 2.4 | 2.45 | 2.1 | 2.1 | 2.2 | 2.2 | 2.4 | 2.2 | 2.34 | 2.1438 | 2.21 | 2 | 2.4 | 2.75 | 2.36 | 3.162 | [6 empty] | 16
| 1 | 4 | 2.3 | 2.37 | 2.3 | 2.2 | 2.2 | 2.6 | 2.3 | 2.3 | 2.39 | 2.2368 | 1.97 | 1.9 | 2.6 | 2.76 | 2.13 | 3.076 | [6 empty] | 2.2664
| 1 | 5 | 2.2 | 2.21 | 2.3 | 2.3 | 2.2 | 2.5 | 2.4 | 2.3 | 2.29 | 2.1713 | 2.16 | 2.3 | 2.6 | 2.77 | 2.16 | 3.172 | [6 empty] | 33
| 1 | 6 | 2.3 | 2.15 | 2.3 | 2.2 | 2.2 | 2.6 | 2.3 | 2.3 | 2.32 | 2.1404 | 2.23 | 2.2 | 2.8 | 2.9 | 2.14 | 3.122
| Mean |  | 2.3 | 2.2667 | 2.2333 | 2.15 | 2.2167 | 2.4667 | 2.3667 | 2.2833 | 2.36 | 2.159 | 2.1283 | 2.1167 | 2.5333 | 2.7833 | 2.1867 | 3.1152
| Median |  | 2.3 | 2.245 | 2.25 | 2.2 | 2.2 | 2.5 | 2.4 | 2.3 | 2.365 | 2.1575 | 2.155 | 2.1 | 2.6 | 2.765 | 2.15 | 3.105
| Std Dev. |  | 0.063246 | 0.12437 | 0.08165 | 0.13784 | 0.040825 | 0.15055 | 0.05164 | 0.040825 | 0.052154 | 0.052836 | 0.099683 | 0.17224 | 0.20656 | 0.058878 | 0.088919 | 0.044029
| Rel.Std.Dev. |  | 2.75% | 5.487% | 3.656% | 6.411% | 1.842% | 6.104% | 2.182% | 1.788% | 2.21% | 2.447% | 4.684% | 8.137% | 8.154% | 2.115% | 4.066% | 1.413%
| PDM3 |  | 1.483% | 0.01264% | -1.458% | -5.135% | -2.194% | 8.837% | 4.425% | 0.748% | 4.131% | -4.738% | -6.091% | -6.606% | 11.78% | 22.81% | -3.517% | 37.45%
| Z-Score (Absolute) |  | 0.16 | 0.05 | 0.27 | 0.81 | 0.38 | 1.24 | 0.6 | 0.05 | 0.55 | 0.75 | 0.95 | 1.03 | 1.68 | 3.3 | 0.57 | 5.46
| Analytical results for Hg in OREAS 605b (Indicative Value < 2 ppm) | [23 empty] | Indicative
Hg | Round | Replicate | 02: Lab | [6 empty] | 1
| No | No | 02: 24 | [6 empty] | ppm
|  |  | 02: 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 02: < 2 | [6 empty] | 1
| 1 | 2 | 02: < 2 | [6 empty] | 7
| 1 | 3 | 02: < 2 | [6 empty] | 16
| 1 | 4 | 02: < 2 | [6 empty] | < 2
| 1 | 5 | 02: < 2 | [6 empty] | 13
| 1 | 6 | 02: < 2
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for Ho in OREAS 605b (Certified Value 0.22 ppm) | [23 empty] | Certified
Ho | Round | Replicate | 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | [6 empty] | 1
| No | No | 02: 01; 03: 03; 05: 11; 06: 14; 07: 16; 08: 23 | [6 empty] | ppm
|  |  | 02: 4A*MS; 03: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 5A*MS | [6 empty] | 2
2
| 1 | 1 | 02: 0.39183; 03: 0.3; 05: 0.23; 06: 0.2; 07: 0.22; 08: 0.217 | [6 empty] | 1
| 1 | 2 | 02: 0.37222; 03: 0.2; 05: 0.24; 06: 0.2; 07: 0.2; 08: 0.216 | [6 empty] | 10
| 1 | 3 | 02: 0.38104; 03: 0.2; 05: 0.23; 06: 0.2; 07: 0.22; 08: 0.217 | [6 empty] | 16
| 1 | 4 | 02: 0.35957; 03: 0.2; 05: 0.24; 06: 0.2; 07: 0.22; 08: 0.206 | [6 empty] | 0.22179
| 1 | 5 | 02: 0.35175; 03: 0.2; 05: 0.24; 06: 0.2; 07: 0.24; 08: 0.216 | [6 empty] | 34
| 1 | 6 | 02: 0.38658; 03: 0.2; 05: 0.24; 06: 0.3; 07: 0.24; 08: 0.212
| Mean |  | 02: 0.37383; 03: 0.21667; 05: 0.23667; 06: 0.21667; 07: 0.22333; 08: 0.214
| Median |  | 02: 0.37663; 03: 0.2; 05: 0.24; 06: 0.2; 07: 0.22; 08: 0.216
| Std Dev. |  | 02: 0.015697; 03: 0.040825; 05: 0.005164; 06: 0.040825; 07: 0.015055; 08: 0.0043359
| Rel.Std.Dev. |  | 02: 4.199%; 03: 18.84%; 05: 2.182%; 06: 18.84%; 07: 6.741%; 08: 2.026%
| PDM3 |  | 02: 68.56%; 03: -2.309%; 05: 6.709%; 06: -2.309%; 07: 0.6974%; 08: -3.511%
| Z-Score (Absolute) |  | 02: 22.23; 03: 0.48; 05: 2.41; 06: 0.48; 07: 0.48; 08: 0.87
| Analytical results for In in OREAS 605b (Certified Value 3.92 ppm) | [23 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | [5 empty] | 2
3
| 1 | 1 | 4.414 | 3.69 | 4.46 | 4.06 | 3.86 | 3.84 | 4.04 | 4.11 | 3.84 | 3.78 | 4.0712 | 3.77 | 3.63 | 3.74 | 4.39 | 3.6 | 4.1078 | [5 empty] | 1
| 1 | 2 | 4.4832 | 3.61 | 3.98 | 4.14 | 3.82 | 3.79 | 4.2 | 4.08 | 3.89 | 3.63 | 4.0431 | 3.37 | 3.66 | 3.4 | 4.33 | 3.56 | 4.1227 | [5 empty] | 27
| 1 | 3 | 4.5228 | 3.66 | 4.16 | 4.09 | 3.75 | 3.81 | 3.95 | 4 | 3.86 | 3.78 | 4.0624 | 3.66 | 3.67 | 3.74 | 4.2 | 3.56 | 4.1619 | [5 empty] | 16
| 1 | 4 | 4.374 | 3.52 | 4.03 | 3.99 | 3.94 | 4 | 4.21 | 3.8 | 3.94 | 3.92 | 3.9973 | 3.33 | 3.59 | 3.8 | 4.17 | 3.6 | 4.0641 | [5 empty] | 3.9156
| 1 | 5 | 4.2597 | 3.52 | 4.07 | 4.2 | 3.89 | 3.89 | 4.15 | 4.03 | 3.77 | 3.88 | 4.0279 | 3.11 | 3.52 | 3.88 | 4.27 | 3.55 | 4.2711 | [5 empty] | 35
| 1 | 6 | 4.3505 | 3.83 | 3.95 | 4.25 | 3.8 | 3.98 | 4.13 | 3.84 | 3.81 | 3.89 | 4.0227 | 3.44 | 3.72 | 3.96 | 4.35 | 3.61 | 4.1676
| Mean |  | 4.4007 | 3.6383 | 4.1083 | 4.1217 | 3.8433 | 3.885 | 4.1133 | 3.9767 | 3.8517 | 3.8133 | 4.0374 | 3.4467 | 3.6317 | 3.7533 | 4.285 | 3.58 | 4.1492
| Median |  | 4.394 | 3.635 | 4.05 | 4.115 | 3.84 | 3.865 | 4.14 | 4.015 | 3.85 | 3.83 | 4.0355 | 3.405 | 3.645 | 3.77 | 4.3 | 3.58 | 4.1423
| Std Dev. |  | 0.094867 | 0.1172 | 0.18734 | 0.094956 | 0.067725 | 0.088261 | 0.10053 | 0.12785 | 0.059805 | 0.10727 | 0.027252 | 0.23788 | 0.06969 | 0.19294 | 0.087121 | 0.026077 | 0.07073
| Rel.Std.Dev. |  | 2.156% | 3.221% | 4.56% | 2.304% | 1.762% | 2.272% | 2.444% | 3.215% | 1.553% | 2.813% | 0.675% | 6.902% | 1.919% | 5.141% | 2.033% | 0.7284% | 1.705%
| PDM3 |  | 12.39% | -7.081% | 4.922% | 5.262% | -1.846% | -0.7817% | 5.05% | 1.559% | -1.633% | -2.612% | 3.112% | -11.98% | -7.252% | -4.144% | 9.434% | -8.571% | 5.966%
| Z-Score (Absolute) |  | 1.52 | 0.73 | 0.66 | 0.7 | 0.12 | 0 | 0.67 | 0.27 | 0.1 | 0.21 | 0.45 | 1.29 | 0.75 | 0.39 | 1.18 | 0.9 | 0.78
| Analytical results for Ir in OREAS 605b (Indicative Value < 5 ppb) | [23 empty] | Indicative
Ir | Round | Replicate | 01: Lab | [5 empty] | 1
| No | No | 01: 23 | [5 empty] | ppb
|  |  | 01: 5A*MS | [5 empty] | 2
2
| 1 | 1 | 01: < 5 | [5 empty] | 1
| 1 | 2 | 01: < 5 | [5 empty] | 8
| 1 | 3 | 01: < 5 | [5 empty] | 16
| 1 | 4 | 01: < 5 | [5 empty] | < 5
| 1 | 5 | 01: < 5 | [5 empty] | 14
| 1 | 6 | 01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for K in OREAS 605b (Certified Value 1.2 wt.%) | [23 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 1.1715 | 1.2196 | 1.24 | 1.22 | 1.21 | 1.14 | 1.19 | 1.19 | 1.18 | 1.39 | 1.2014 | 1.1942 | 1.196 | 1.19 | 1.19 | 1.33 | 1.19 | 1.12 | 1.1853 | 1.2293 | 1.219 | 1.191 | 1
| 1 | 2 | 1.1579 | 1.223 | 1.24 | 1.22 | 1.21 | 1.17 | 1.18 | 1.19 | 1.17 | 1.32 | 1.1825 | 1.1475 | 1.2117 | 1.2 | 1.19 | 1.33 | 1.18 | 1.17 | 1.1875 | 1.2216 | 1.203 | 1.159 | #N/A
| 1 | 3 | 1.1314 | 1.2127 | 1.24 | 1.22 | 1.21 | 1.17 | 1.16 | 1.19 | 1.19 | 1.33 | 1.1863 | 1.1629 | 1.2182 | 1.18 | 1.19 | 1.31 | 1.19 | 1.21 | 1.1946 | 1.2155 | 1.241 | 1.201 | 16
| 1 | 4 | 1.1693 | 1.2098 | 1.22 | 1.23 | 1.19 | 1.16 | 1.18 | 1.26 | 1.16 | 1.32 | 1.2299 | 1.1535 | 1.2051 | 1.21 | 1.21 | 1.38 | 1.19 | 1.17 | 1.1317 | 1.2408 | 1.228 | 1.152 | 1.1958
| 1 | 5 | 1.1107 | 1.1508 | 1.27 | 1.23 | 1.21 | 1.17 | 1.17 | 1.28 | 1.15 | 1.34 | 1.2298 | 1.1893 | 1.2295 | 1.14 | 1.21 | 1.39 | 1.2 | 1.15 | 1.1833 | 1.2405 | 1.201 | 1.199 | 36
| 1 | 6 | 1.1776 | 1.2206 | 1.18 | 1.24 | 1.22 | 1.21 | 1.2 | 1.27 | 1.15 | 1.31 | 1.219 | 1.1615 | 1.2261 | 1.13 | 1.2 | 1.34 | 1.18 | 1.16 | 1.1916 | 1.2018 | 1.219 | 1.148
| Mean |  | 1.1531 | 1.2061 | 1.2317 | 1.2267 | 1.2083 | 1.17 | 1.18 | 1.23 | 1.1667 | 1.335 | 1.2082 | 1.1681 | 1.2144 | 1.175 | 1.1983 | 1.3467 | 1.1883 | 1.1633 | 1.179 | 1.2249 | 1.2185 | 1.175
| Median |  | 1.1636 | 1.2162 | 1.24 | 1.225 | 1.21 | 1.17 | 1.18 | 1.225 | 1.165 | 1.325 | 1.2102 | 1.1622 | 1.2149 | 1.185 | 1.195 | 1.335 | 1.19 | 1.165 | 1.1864 | 1.2255 | 1.219 | 1.175
| Std Dev. |  | 0.026431 | 0.027544 | 0.029944 | 0.008165 | 0.0098319 | 0.022804 | 0.014142 | 0.044272 | 0.01633 | 0.02881 | 0.021168 | 0.019185 | 0.012741 | 0.032711 | 0.0098319 | 0.031411 | 0.0075277 | 0.029439 | 0.023519 | 0.015153 | 0.015123 | 0.024585
| Rel.Std.Dev. |  | 2.292% | 2.284% | 2.431% | 0.6656% | 0.8137% | 1.949% | 1.198% | 3.599% | 1.4% | 2.158% | 1.752% | 1.642% | 1.049% | 2.784% | 0.8205% | 2.333% | 0.6335% | 2.531% | 1.995% | 1.237% | 1.241% | 2.092%
| PDM3 |  | -3.576% | 0.858% | 2.997% | 2.579% | 1.046% | -2.159% | -1.323% | 2.858% | -2.438% | 11.64% | 1.031% | -2.315% | 1.556% | -1.741% | 0.2099% | 12.61% | -0.6264% | -2.717% | -1.406% | 2.433% | 1.896% | -1.741%
| Z-Score (Absolute) |  | 1.28 | 0.1 | 0.77 | 0.64 | 0.16 | 0.84 | 0.58 | 0.73 | 0.93 | 3.47 | 0.16 | 0.89 | 0.32 | 0.71 | 0.1 | 3.77 | 0.36 | 1.01 | 0.61 | 0.59 | 0.43 | 0.71
| Analytical results for La in OREAS 605b (Indicative Value 9.46 ppm) | [23 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
2
| 1 | 1 | 11.268 | 3.2 | < 5 | 5.6 | 9.2 | 9 | 6.9 | 10.4 | 19.48 | 3 | 5.8 | 13.7 | 10.4 | 4.24 | 3.8 | 12.697 | 14.0146 | 15.11 | 7.935 |  |  |  | 1
| 1 | 2 | 11.765 | 2.5 | < 5 | 6.2 | 10 | 8.5 | 7.4 | 10 | 19.5 | 4 | 6.7 | 13 | 10.6 | 4.46 | 3.7 | 12.419 | 14.185 | 15.94 | 7.548 |  |  |  | 5
| 1 | 3 | 11.3814 | 2.7 | < 5 | 5.1 | 9.9 | 9.5 | 3.8 | 10.3 | 17.87 | 4 | 5.7 | 12.6 | 10 | 4.79 | 2.9 | 13.216 | 14.2069 | 14.68 | 7.317 |  |  |  | 16
| 1 | 4 | 11.665 | 2.7 | < 5 | 5.7 | 9.3 | 9.1 | 9 | 9.9 | 18.16 | 5 | 5.9 | 12.7 | 10.4 | 4.61 | 3.7 | 12.666 | 13.649 | 15.09 | 8.055 |  |  |  | 9.4558
| 1 | 5 | 11.028 | 3 | < 5 | 5.9 | 9.4 | 8.7 | 9.2 | 10 | 17.57 | 5 | 6.4 | 11.4 | 10.7 | 4.97 | 3.9 | 12.952 | 13.6053 | 15.13 | 6.864 |  |  |  | 11
| 1 | 6 | 11.369 | 2.5 | < 5 | 6.4 | 11.7 | 8.9 | 8.3 | 9.4 | 17.48 | 5 | 6.9 | 12.1 | 10.8 | 4.54 | 3.2 | 12.655 | 13.841 | 14.85 | 6.53
| Mean |  | 11.4127 | 2.7667 |  | 5.8167 | 9.9167 | 8.95 | 7.4333 | 10 | 18.3433 | 4.3333 | 6.2333 | 12.5833 | 10.4833 | 4.6017 | 3.5333 | 12.7675 | 13.9169 | 15.1333 | 7.3748
| Median |  | 11.3752 | 2.7 |  | 5.8 | 9.65 | 8.95 | 7.85 | 10 | 18.015 | 4.5 | 6.15 | 12.65 | 10.5 | 4.575 | 3.7 | 12.6815 | 13.9278 | 15.1 | 7.4325
| Std Dev. |  | 0.26818 | 0.28048 |  | 0.46224 | 0.93256 | 0.34496 | 1.9906 | 0.35214 | 0.91988 | 0.8165 | 0.50465 | 0.78337 | 0.28577 | 0.2553 | 0.39328 | 0.27733 | 0.26083 | 0.43334 | 0.5974
| Rel.Std.Dev. |  | 2.35% | 10.14% |  | 7.947% | 9.404% | 3.854% | 26.78% | 3.521% | 5.015% | 18.84% | 8.096% | 6.225% | 2.726% | 5.548% | 11.13% | 2.172% | 1.874% | 2.864% | 8.101%
| PDM3 |  | 20.7% | -70.74% |  | -38.49% | 4.874% | -5.349% | -21.39% | 5.755% | 93.99% | -54.17% | -34.08% | 33.07% | 10.87% | -51.34% | -62.63% | 35.02% | 47.18% | 60.04% | -22.01%
| Z-Score (Absolute) |  | 0.41 | 1.38 | 1.43 | 0.75 | 0.1 | 0.1 | 0.41 | 0.12 | 1.84 | NA | 0.66 | 0.65 | 0.22 | 1 | 1.22 | 0.69 | 0.93 | 1.18 | 0.42
Results from laboratory 14 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 605b (Certified Value 28.9 ppm) | [23 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 26.7415 | 28.7 | 27.3 | 29 | 30.4 | 27.6 | 28.3 | 32.1 | 24.7 | 28.8 | 27.5 | 25.6689 | 28 | 27.5 | 29 | 31 | 33.4 | 29 | 30.194 | 28.7309 | 31.34 | 27.724 | 1
| 1 | 2 | 26.2688 | 28.2 | 28.1 | 29 | 31.1 | 28.5 | 27.6 | 32.8 | 24.7 | 28 | 27.5 | 25.5772 | 31 | 27.7 | 25.5 | 31 | 34.5 | 30 | 29.787 | 28.4254 | 31.98 | 27.772 | #N/A
| 1 | 3 | 26.8369 | 28.2 | 30.5 | 28 | 30.7 | 28.4 | 27.5 | 28.5 | 23.8 | 27.6 | 27.7 | 25.7608 | 31 | 27.4 | 30.5 | 30 | 35.2 | 31 | 30.281 | 28.2383 | 29.03 | 28.981 | 16
| 1 | 4 | 26.6058 | 28.5 | 30.6 | 28 | 28.5 | 31.1 | 28.3 | 27.1 | 22.9 | 28.6 | 27.8 | 25.8541 | 31 | 28 | 30 | 33 | 31.9 | 29 | 28.584 | 29.3294 | 29.77 | 27.437 | 28.8883
| 1 | 5 | 26.1079 | 27.6 | 30.6 | 28 | 28.7 | 29.9 | 28.1 | 25 | 23.6 | 28.6 | 28.2 | 26.1818 | 32 | 26.5 | 29.5 | 31 | 35.5 | 28 | 30.041 | 28.9974 | 30.36 | 29.397 | 37
| 1 | 6 | 26.6938 | 28.9 | 31 | 29 | 28.7 | 32.2 | 28.1 | 20.1 | 22.4 | 27.5 | 27.6 | 26.1726 | 30 | 25.8 | 30 | 29 | 32.6 | 30 | 30.552 | 28.1326 | 30.52 | 29.477
| Mean |  | 26.5425 | 28.35 | 29.6833 | 28.5 | 29.6833 | 29.6167 | 27.9833 | 27.6 | 23.6833 | 28.1833 | 27.7167 | 25.8692 | 30.5 | 27.15 | 29.0833 | 30.8333 | 33.85 | 29.5 | 29.9065 | 28.6423 | 30.5 | 28.4647
| Median |  | 26.6498 | 28.35 | 30.55 | 28.5 | 29.55 | 29.2 | 28.1 | 27.8 | 23.7 | 28.3 | 27.65 | 25.8075 | 31 | 27.45 | 29.75 | 31 | 33.95 | 29.5 | 30.1175 | 28.5782 | 30.44 | 28.3765
| Std Dev. |  | 0.28878 | 0.45935 | 1.5664 | 0.54772 | 1.1737 | 1.7747 | 0.34881 | 4.7193 | 0.93256 | 0.56006 | 0.26394 | 0.25579 | 1.3784 | 0.83126 | 1.828 | 1.3292 | 1.4516 | 1.0488 | 0.69584 | 0.46393 | 1.0592 | 0.92143
| Rel.Std.Dev. |  | 1.088% | 1.62% | 5.277% | 1.922% | 3.954% | 5.992% | 1.246% | 17.1% | 3.938% | 1.987% | 0.9523% | 0.9888% | 4.519% | 3.062% | 6.285% | 4.311% | 4.288% | 3.555% | 2.327% | 1.62% | 3.473% | 3.237%
| PDM3 |  | -8.12% | -1.863% | 2.752% | -1.344% | 2.752% | 2.521% | -3.133% | -4.46% | -18.02% | -2.44% | -4.056% | -10.45% | 5.579% | -6.017% | 0.6751% | 6.733% | 17.18% | 2.117% | 3.525% | -0.8515% | 5.579% | -1.467%
| Z-Score (Absolute) |  | 1.27 | 0.14 | 0.7 | 0.04 | 0.7 | 0.65 | 0.37 | 0.61 | 3.06 | 0.24 | 0.53 | 1.69 | 1.21 | 0.89 | 0.32 | 1.41 | 3.3 | 0.58 | 0.83 | 0.04 | 1.21 | 0.07
| Analytical results for Lu in OREAS 605b (Certified Value 87.6 ppb) | [23 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 03; 02: 11; 03: 14; 04: 16; 05: 20; 06: 21; 07: 23 | ppb
|  |  | 01: 4A*OES/MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 07: 5A*MS | 1
1
| 1 | 1 | 01: < 100; 02: 80; 03: < 100; 04: 100; 05: 100; 06: 77.1; 07: 83 | 1
| 1 | 2 | 01: < 100; 02: 90; 03: < 100; 04: 100; 05: 100; 06: 81.5; 07: 78 | 29
| 1 | 3 | 01: 100; 02: 80; 03: < 100; 04: 120; 05: 100; 06: 81.1; 07: 82 | 16
| 1 | 4 | 01: < 100; 02: 80; 03: < 100; 04: 100; 05: 90; 06: 78.5; 07: 71 | 87.6056
| 1 | 5 | 01: 100; 02: 90; 03: < 100; 04: 100; 05: 100; 06: 81.9; 07: 84 | 38
| 1 | 6 | 01: 100; 02: 90; 03: < 100; 04: 120; 05: 100; 06: 72.9; 07: 82
| Mean |  | 01: 100; 02: 85; 04: 106.6667; 05: 98.3333; 06: 78.8333; 07: 80
| Median |  | 01: 100; 02: 85; 04: 100; 05: 100; 06: 79.8; 07: 82
| Std Dev. |  | 01: 0; 02: 5.4772; 04: 10.328; 05: 4.0825; 06: 3.4633; 07: 4.858
| Rel.Std.Dev. |  | 01: 0%; 02: 6.444%; 04: 9.682%; 05: 4.152%; 06: 4.393%; 07: 6.072%
| PDM3 |  | 01: 14.15%; 02: -2.974%; 04: 21.76%; 05: 12.25%; 06: -10.01%; 07: -8.682%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67; 03: 4.05; 04: 3.6; 05: 2.47; 06: 0.16; 07: 0
| Analytical results for Mg in OREAS 605b (Certified Value 1731 ppm) | [23 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 1600.8 | 1758 | 1700 | 1400 | 1700 | 1500 | 1700 | 1500 | 1700 | 2040 | 1775 | 1737.2493 | 1789 | 1800 | 1800 | 1800 | 1760 | 1800 | 1904.295 | 1866.1218 | 1818 | 1580 | 1
| 1 | 2 | 1633.1 | 1726 | 1800 | 1300 | 1700 | 1500 | 1700 | 1500 | 1600 | 2040 | 1758 | 1675.2024 | 1795 | 1700 | 1800 | 1800 | 1760 | 1800 | 1848.536 | 1852.3687 | 1825 | 1550 | #N/A
| 1 | 3 | 1590.6 | 1720 | 1900 | 1400 | 1700 | 1600 | 1600 | 1600 | 1700 | 2060 | 1771 | 1689.0955 | 1777 | 1700 | 1800 | 1800 | 1810 | 1800 | 1861.931 | 1865.6046 | 1824 | 1620 | 16
| 1 | 4 | 1594 | 1728 | 1900 | 1400 | 1700 | 1500 | 1700 | 1700 | 1600 | 2050 | 1795 | 1692.7542 | 1805 | 1800 | 1800 | 1800 | 1710 | 1700 | 1799.482 | 1882.3605 | 1832 | 1560 | 1731.1584
| 1 | 5 | 1593.5 | 1664 | 1800 | 1300 | 1700 | 1600 | 1700 | 1700 | 1600 | 2060 | 1783 | 1727.3755 | 1824 | 1700 | 1800 | 1900 | 1770 | 1800 | 1885.53 | 1850.6748 | 1820 | 1590 | 39
| 1 | 6 | 1672.7 | 1749 | 1800 | 1400 | 1700 | 1700 | 1700 | 1700 | 1600 | 2040 | 1780 | 1696.7742 | 1701 | 1600 | 1800 | 1700 | 1750 | 1800 | 1932.994 | 1836.2624 | 1899 | 1550
| Mean |  | 1614.1167 | 1724.1667 | 1816.6667 | 1366.6667 | 1700 | 1566.6667 | 1683.3333 | 1616.6667 | 1633.3333 | 2048.3333 | 1777 | 1703.0752 | 1781.8333 | 1716.6667 | 1800 | 1800 | 1760 | 1783.3333 | 1872.128 | 1858.8988 | 1836.3333 | 1575
| Median |  | 1597.4 | 1727 | 1800 | 1400 | 1700 | 1550 | 1700 | 1650 | 1600 | 2045 | 1777.5 | 1694.7642 | 1792 | 1700 | 1800 | 1800 | 1760 | 1800 | 1873.7305 | 1858.9867 | 1824.5 | 1570
| Std Dev. |  | 32.7184 | 32.9145 | 75.2773 | 51.6398 | 0 | 81.6497 | 40.8248 | 98.3192 | 51.6398 | 9.8319 | 12.4097 | 23.9872 | 42.654 | 75.2773 | 0 | 63.2456 | 32.249 | 40.8248 | 46.5983 | 15.9391 | 31.0784 | 27.3861
| Rel.Std.Dev. |  | 2.027% | 1.909% | 4.144% | 3.779% | 0% | 5.212% | 2.425% | 6.082% | 3.162% | 0.48% | 0.6983% | 1.408% | 2.394% | 4.385% | 0% | 3.514% | 1.832% | 2.289% | 2.489% | 0.8574% | 1.692% | 1.739%
| PDM3 |  | -6.761% | -0.4039% | 4.939% | -21.05% | -1.8% | -9.502% | -2.763% | -6.614% | -5.651% | 18.32% | 2.648% | -1.622% | 2.927% | -0.8371% | 3.977% | 3.977% | 1.666% | 3.014% | 8.143% | 7.379% | 6.075% | -9.02%
| Z-Score (Absolute) |  | 1.29 | 0.18 | 0.75 | 3.8 | 0.43 | 1.77 | 0.59 | 1.27 | 1.1 | 3.1 | 0.35 | 0.39 | 0.4 | 0.26 | 0.59 | 0.59 | 0.18 | 0.42 | 1.32 | 1.18 | 0.95 | 1.69
| Analytical results for Mn in OREAS 605b (Certified Value 109 ppm) | [23 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 111.98 | 111 | 118 | 138 | 108 | 105 | 111 | 115 | 111 | 110 | 112 | 105.6215 | 114 | 105 | 110 | 132 | 111 | 105 | 113.436 | 110.7298 | 102.5 | 108.273 | 1
| 1 | 2 | 107.95 | 109 | 114 | 137 | 105 | 103 | 110 | 116 | 110 | 116 | 110 | 105.0016 | 115 | 105 | 112 | 126 | 113 | 107 | 106.955 | 110.3597 | 101.9 | 100.643 | 30
| 1 | 3 | 103.03 | 108 | 117 | 137 | 106 | 107 | 112 | 110 | 111 | 115 | 110 | 104.3602 | 116 | 105 | 104 | 127 | 115 | 112 | 108.587 | 110.9812 | 101.5 | 113.37 | 16
| 1 | 4 | 110.96 | 111 | 116 | 139 | 105 | 105 | 108 | 112 | 107 | 117 | 115 | 103.6518 | 115 | 109 | 112 | 125 | 107 | 105 | 102.972 | 111.4986 | 101.8 | 106.032 | 109.2601
| 1 | 5 | 105.39 | 105 | 118 | 138 | 107 | 106 | 110 | 112 | 109 | 116 | 113 | 104.9838 | 117 | 100 | 108 | 128 | 112 | 104 | 108.64 | 110.095 | 103.3 | 105.385 | 40
| 1 | 6 | 115.09 | 110 | 113 | 138 | 110 | 115 | 113 | 112 | 107 | 114 | 115 | 103.7475 | 112 | 99 | 110 | 130 | 108 | 110 | 111.691 | 110.6902 | 102.7 | 101.613
| Mean |  | 109.0667 | 109 | 116 | 137.8333 | 106.8333 | 106.8333 | 110.6667 | 112.8333 | 109.1667 | 114.6667 | 112.5 | 104.5611 | 114.8333 | 103.8333 | 109.3333 | 128 | 111 | 107.1667 | 108.7135 | 110.7257 | 102.2833 | 105.886
| Median |  | 109.455 | 109.5 | 116.5 | 138 | 106.5 | 105.5 | 110.5 | 112 | 109.5 | 115.5 | 112.5 | 104.672 | 115 | 105 | 110 | 127.5 | 111.5 | 106 | 108.6135 | 110.71 | 102.2 | 105.7085
| Std Dev. |  | 4.4623 | 2.2804 | 2.0976 | 0.75277 | 1.9408 | 4.2151 | 1.7512 | 2.2286 | 1.8348 | 2.5033 | 2.2583 | 0.778 | 1.7224 | 3.7103 | 3.0111 | 2.6077 | 3.0332 | 3.1885 | 3.6673 | 0.48902 | 0.67057 | 4.6417
| Rel.Std.Dev. |  | 4.091% | 2.092% | 1.808% | 0.5461% | 1.817% | 3.945% | 1.582% | 1.975% | 1.681% | 2.183% | 2.007% | 0.7441% | 1.5% | 3.573% | 2.754% | 2.037% | 2.733% | 2.975% | 3.373% | 0.4416% | 0.6556% | 4.384%
| PDM3 |  | -0.1771% | -0.2381% | 6.169% | 26.15% | -2.221% | -2.221% | 1.287% | 3.27% | -0.08556% | 4.948% | 2.965% | -4.301% | 5.101% | -4.967% | 0.06698% | 17.15% | 1.592% | -1.916% | -0.5003% | 1.341% | -6.386% | -3.088%
| Z-Score (Absolute) |  | 0.04 | 0.06 | 1.61 | 6.8 | 0.58 | 0.58 | 0.34 | 0.85 | 0.02 | 1.29 | 0.77 | 1.12 | 1.33 | 1.29 | 0.02 | 4.46 | 0.42 | 0.5 | 0.13 | 0.35 | 1.66 | 0.8
| Analytical results for Mo in OREAS 605b (Certified Value 17.3 ppm) | [23 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 20.0403 | 16.8 | 18.76 | 17 | 16.85 | 16.55 | 18 | 18.85 | 18.35 | 17.1 | 17.5 | 17.5793 | 16.48 | 18.2 | 20 | 17.49 | 16.6 | 15 | 17.3733 | 15.5616 | 16.26 | 15.748 | 1
| 1 | 2 | 20.127 | 16.5 | 18.16 | 18 | 17.5 | 16.8 | 17.75 | 19.5 | 19.35 | 17 | 17.3 | 17.7287 | 17.18 | 18.6 | 18 | 17.29 | 16.8 | 15 | 17.0709 | 17.3572 | 16.18 | 15.027 | 31
| 1 | 3 | 19.98 | 16.8 | 19.04 | 17 | 17.35 | 17.15 | 17.75 | 18.05 | 17.9 | 17.1 | 17.3 | 17.1714 | 16.19 | 18.9 | 20 | 18.01 | 16.8 | 16 | 17.4709 | 17.2882 | 16.64 | 15.405 | 16
| 1 | 4 | 19.907 | 16.6 | 19.46 | 17 | 17.2 | 18 | 18.6 | 18.5 | 17 | 17.3 | 17.5 | 17.5908 | 16.82 | 18.5 | 19.5 | 17.35 | 16.8 | 16 | 16.7296 | 20.9813 | 15.93 | 15.201 | 17.3099
| 1 | 5 | 20.1703 | 16.1 | 18.66 | 17 | 17.55 | 16.7 | 17.95 | 18.65 | 18.05 | 17.2 | 17.5 | 17.6077 | 17.69 | 17.9 | 20 | 17.5 | 17 | 16 | 17.4919 | 16.1177 | 16.14 | 15.491 | 41
| 1 | 6 | 20.0981 | 16.7 | 18.15 | 17 | 17.4 | 17.35 | 17.9 | 17.85 | 17 | 17 | 17.8 | 17.258 | 16.8 | 18.7 | 20 | 17.97 | 16.8 | 15 | 17.5854 | 17.2253 | 15.94 | 15.552
| Mean |  | 20.0538 | 16.5833 | 18.705 | 17.1667 | 17.3083 | 17.0917 | 17.9917 | 18.5667 | 17.9417 | 17.1167 | 17.4833 | 17.4893 | 16.86 | 18.4667 | 19.5833 | 17.6017 | 16.8 | 15.5 | 17.287 | 17.4219 | 16.1817 | 15.404
| Median |  | 20.0692 | 16.65 | 18.71 | 17 | 17.375 | 16.975 | 17.925 | 18.575 | 17.975 | 17.1 | 17.5 | 17.585 | 16.81 | 18.55 | 20 | 17.495 | 16.8 | 15.5 | 17.4221 | 17.2567 | 16.16 | 15.448
| Std Dev. |  | 0.098043 | 0.26394 | 0.5084 | 0.40825 | 0.25577 | 0.53424 | 0.3153 | 0.59048 | 0.88737 | 0.1169 | 0.18348 | 0.22106 | 0.52722 | 0.36148 | 0.80104 | 0.31167 | 0.12649 | 0.54772 | 0.32548 | 1.8917 | 0.26065 | 0.25735
| Rel.Std.Dev. |  | 0.4889% | 1.592% | 2.718% | 2.378% | 1.478% | 3.126% | 1.753% | 3.18% | 4.946% | 0.683% | 1.049% | 1.264% | 3.127% | 1.957% | 4.09% | 1.771% | 0.7529% | 3.534% | 1.883% | 10.86% | 1.611% | 1.671%
| PDM3 |  | 15.85% | -4.197% | 8.059% | -0.8276% | -0.009142% | -1.261% | 3.938% | 7.26% | 3.65% | -1.116% | 1.002% | 1.036% | -2.599% | 6.683% | 13.13% | 1.685% | -2.946% | -10.46% | -0.1322% | 0.6469% | -6.518% | -11.01%
| Z-Score (Absolute) |  | 3.1 | 1.04 | 1.49 | NA | 0.17 | 0.43 | 0.64 | 1.33 | 0.58 | 0.4 | 0.04 | 0.04 | 0.71 | 1.21 | 2.54 | 0.18 | 0.78 | NA | 0.2 | 0.04 | 1.51 | 2.44
Results from laboratories 4 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 605b (Certified Value 0.527 wt.%) | [23 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.49433 | 0.5464 | 0.477 | 0.46 | 0.52 | 0.5 | 0.52 | 0.52 | 0.53 | 0.607 | 0.5106 | 0.51588 | 0.4916 | 0.5 | 0.53 | 0.72 | 0.534 | 0.51 | 0.57065 | 0.66714 | 0.5818 | 0.52 | 1
| 1 | 2 | 0.49167 | 0.5458 | 0.545 | 0.46 | 0.53 | 0.51 | 0.52 | 0.51 | 0.52 | 0.607 | 0.5045 | 0.50757 | 0.5424 | 0.51 | 0.53 | 0.7 | 0.536 | 0.52 | 0.56235 | 0.64924 | 0.5802 | 0.495 | #N/A
| 1 | 3 | 0.49028 | 0.539 | 0.512 | 0.46 | 0.53 | 0.51 | 0.51 | 0.52 | 0.53 | 0.607 | 0.5097 | 0.51732 | 0.5443 | 0.52 | 0.53 | 0.68 | 0.54 | 0.54 | 0.5591 | 0.61861 | 0.5871 | 0.505 | 16
| 1 | 4 | 0.49387 | 0.5405 | 0.524 | 0.47 | 0.56 | 0.51 | 0.51 | 0.55 | 0.52 | 0.61 | 0.5176 | 0.50369 | 0.5349 | 0.53 | 0.54 | 0.74 | 0.53 | 0.52 | 0.53883 | 0.68674 | 0.5821 | 0.504 | 0.52662
| 1 | 5 | 0.49016 | 0.5145 | 0.524 | 0.47 | 0.55 | 0.51 | 0.51 | 0.56 | 0.52 | 0.603 | 0.5117 | 0.52202 | 0.549 | 0.49 | 0.53 | 0.72 | 0.525 | 0.51 | 0.57664 | 0.66773 | 0.5789 | 0.516 | 42
| 1 | 6 | 0.49474 | 0.5423 | 0.533 | 0.47 | 0.56 | 0.52 | 0.52 | 0.55 | 0.51 | 0.599 | 0.5133 | 0.51282 | 0.502 | 0.48 | 0.53 | 0.7 | 0.532 | 0.52 | 0.57058 | 0.59716 | 0.5754 | 0.506
| Mean |  | 0.49251 | 0.53808 | 0.51917 | 0.465 | 0.54167 | 0.51 | 0.515 | 0.535 | 0.52167 | 0.6055 | 0.51123 | 0.51322 | 0.52737 | 0.505 | 0.53167 | 0.71 | 0.53283 | 0.52 | 0.56303 | 0.64777 | 0.58092 | 0.50767
| Median |  | 0.49277 | 0.5414 | 0.524 | 0.465 | 0.54 | 0.51 | 0.515 | 0.535 | 0.52 | 0.607 | 0.51115 | 0.51435 | 0.53865 | 0.505 | 0.53 | 0.71 | 0.533 | 0.52 | 0.56647 | 0.65819 | 0.581 | 0.5055
| Std Dev. |  | 0.0020657 | 0.011911 | 0.023387 | 0.0054772 | 0.017224 | 0.0063246 | 0.0054772 | 0.020736 | 0.0075277 | 0.0038859 | 0.0043154 | 0.0066974 | 0.024332 | 0.018708 | 0.0040825 | 0.020976 | 0.0051543 | 0.010954 | 0.013429 | 0.033782 | 0.003884 | 0.0090037
| Rel.Std.Dev. |  | 0.4194% | 2.214% | 4.505% | 1.178% | 3.18% | 1.24% | 1.064% | 3.876% | 1.443% | 0.6418% | 0.8441% | 1.305% | 4.614% | 3.705% | 0.7679% | 2.954% | 0.9673% | 2.107% | 2.385% | 5.215% | 0.6686% | 1.774%
| PDM3 |  | -6.478% | 2.177% | -1.415% | -11.7% | 2.857% | -3.156% | -2.207% | 1.591% | -0.9407% | 14.98% | -2.922% | -2.545% | 0.1417% | -4.106% | 0.9582% | 34.82% | 1.18% | -1.257% | 6.913% | 23% | 10.31% | -3.599%
| Z-Score (Absolute) |  | 1.54 | 0.65 | 0.26 | 2.86 | 0.82 | 0.7 | 0.46 | 0.5 | 0.14 | 3.89 | 0.64 | 0.54 | 0.14 | 0.94 | 0.34 | 8.91 | 0.4 | 0.22 | 1.85 | 5.92 | 2.71 | 0.81
| Analytical results for Nb in OREAS 605b (Certified Value 6.8 ppm) | [23 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 7.2865 | 7.1 | 6.8 | 5 | 6 | 6.3 | 6.7 | 7 | 7 | 6.1 | 6.93 | 6.1885 | 6.88 | 7.1 | 6 | 7.7 | 6.95 | 6.613 | 5.4618 |  |  |  | 1
| 1 | 2 | 7.1502 | 6.9 | 6.7 | 5 | 6.2 | 6.5 | 6.6 | 7.3 | 6.9 | 5.6 | 6.92 | 6.1733 | 6.99 | 6.9 | 5 | 7.3 | 6.39 | 6.642 | 5.2849 |  |  |  | 32
| 1 | 3 | 7.0574 | 7.1 | 7.07 | 5 | 6.1 | 6.7 | 6.8 | 6.8 | 6.8 | 5.6 | 6.73 | 6.3104 | 6.35 | 6.8 | 6 | 7.4 | 6.63 | 6.692 | 5.6233 |  |  |  | 16
| 1 | 4 | 7.1401 | 6.8 | 7.48 | 5 | 7 | 6.8 | 6.9 | 7.7 | 6.5 | 5.8 | 6.88 | 6.1543 | 7.48 | 6.9 | 5.5 | 7.2 | 6.82 | 6.597 | 5.2772 |  |  |  | 6.7952
| 1 | 5 | 7.0367 | 6.7 | 6.83 | 5 | 7 | 6.8 | 6.9 | 7.6 | 6.8 | 5.8 | 6.97 | 6.2828 | 7.02 | 6.7 | 6 | 7.3 | 6.89 | 6.894 | 5.4685 |  |  |  | 43
| 1 | 6 | 7.0087 | 7.1 | 6.48 | 5 | 7 | 6.9 | 6.8 | 7.3 | 6.3 | 5.7 | 6.87 | 6.3104 | 5.74 | 6.8 | 5.5 | 7.4 | 6.85 | 6.548 | 5.4775
| Mean |  | 7.1133 | 6.95 | 6.8933 | 5 | 6.55 | 6.6667 | 6.7833 | 7.2833 | 6.7167 | 5.7667 | 6.8833 | 6.2366 | 6.7433 | 6.8667 | 5.6667 | 7.3833 | 6.755 | 6.6643 | 5.4322
| Median |  | 7.0988 | 7 | 6.815 | 5 | 6.6 | 6.75 | 6.8 | 7.3 | 6.8 | 5.75 | 6.9 | 6.2357 | 6.935 | 6.85 | 5.75 | 7.35 | 6.835 | 6.6275 | 5.4651
| Std Dev. |  | 0.102 | 0.17607 | 0.34535 | 0 | 0.49699 | 0.22509 | 0.1169 | 0.34303 | 0.26394 | 0.18619 | 0.083347 | 0.072267 | 0.60994 | 0.13663 | 0.40825 | 0.17224 | 0.20897 | 0.12224 | 0.13152
| Rel.Std.Dev. |  | 1.434% | 2.533% | 5.01% | 0% | 7.588% | 3.376% | 1.723% | 4.71% | 3.93% | 3.229% | 1.211% | 1.159% | 9.045% | 1.99% | 7.204% | 2.333% | 3.094% | 1.834% | 2.421%
| PDM3 |  | 4.681% | 2.278% | 1.444% | -26.42% | -3.609% | -1.892% | -0.1748% | 7.183% | -1.156% | -15.14% | 1.297% | -8.22% | -0.7634% | 1.052% | -16.61% | 8.655% | -0.5917% | -1.926% | -20.06%
| Z-Score (Absolute) |  | 1.43 | 0.79 | 0.57 | NA | 0.78 | 0.32 | 0.13 | 2.1 | 0.13 | 3.86 | 0.53 | 2.01 | 0.02 | 0.46 | 4.25 | 2.49 | 0.02 | 0.33 | 5.17
Results from laboratory 4 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 605b (Certified Value 10.4 ppm) | [23 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 23 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 5A*MS |  |  |  | 1
1
| 1 | 1 | 01: 12.2197; 02: 6.6; 03: 12.52; 04: 7.7; 05: 11.3; 06: 10.824 |  |  |  | 1
| 1 | 2 | 01: 12.2984; 02: 6.6; 03: 12.79; 04: 7.7; 05: 10.5; 06: 10.829 |  |  |  | 13
| 1 | 3 | 01: 12.6; 02: 7.7; 03: 13.24; 04: 7.4; 05: 11.3; 06: 11.143 |  |  |  | 16
| 1 | 4 | 01: 12.1853; 02: 7.4; 03: 13.35; 04: 9; 05: 10.7; 06: 10.954 |  |  |  | 10.3811
| 1 | 5 | 01: 12.6363; 02: 6.6; 03: 13.45; 04: 8.4; 05: 10.7; 06: 11.211 |  |  |  | 44
| 1 | 6 | 01: 12.2157; 02: 6.4; 03: 13; 04: 8.6; 05: 11.7; 06: 10.934
| Mean |  | 01: 12.3592; 02: 6.8833; 03: 13.0583; 04: 8.1333; 05: 11.0333; 06: 10.9825
| Median |  | 01: 12.2591; 02: 6.6; 03: 13.12; 04: 8.05; 05: 11; 06: 10.944
| Std Dev. |  | 01: 0.20433; 02: 0.53072; 03: 0.35684; 04: 0.62503; 05: 0.46762; 06: 0.16113
| Rel.Std.Dev. |  | 01: 1.653%; 02: 7.71%; 03: 2.733%; 04: 7.685%; 05: 4.238%; 06: 1.467%
| PDM3 |  | 01: 19.05%; 02: -33.69%; 03: 25.79%; 04: -21.65%; 05: 6.283%; 06: 5.793%
| Z-Score (Absolute) |  | 01: 0.54; 02: 1.64; 03: 0.81; 04: 1.14; 05: 0.01; 06: 0.01
| Analytical results for Ni in OREAS 605b (Certified Value 21.9 ppm) | [23 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 22.022 | 22 | 23.3 | 24 | 21.9 | 22.6 | 20.1 | 21.1 | 21.8 | 22 | 21.8 | 21.6383 | 22.5 | 18.8 | 20 | 24.6 | 21.2 | 20 | 20.746 | 21.411 | 22.32 | 22.953 | 1
| 1 | 2 | 21.9703 | 22 | 22.4 | 25 | 22.4 | 22.4 | 20.1 | 21.9 | 21.6 | 22 | 21.7 | 21.6999 | 21.1 | 19.5 | 22 | 23.3 | 21.3 | 20 | 21.188 | 21.8495 | 21.78 | 22.471 | 34
| 1 | 3 | 22.0136 | 22 | 24.1 | 25 | 22.2 | 23 | 20.9 | 22 | 21.1 | 21 | 22 | 22.9178 | 20.7 | 19 | 24 | 23.9 | 21.6 | 20 | 22.002 | 22.9871 | 21.37 | 21.164 | 16
| 1 | 4 | 22.1771 | 22 | 25.1 | 24 | 21.5 | 24 | 20.9 | 22.3 | 19.7 | 21 | 22.7 | 22.1602 | 22.7 | 19.3 | 22 | 23.1 | 20.7 | 20 | 20.2 | 22.4306 | 21.53 | 24.677 | 21.8785
| 1 | 5 | 21.9505 | 21 | 22.9 | 24 | 22.3 | 22.5 | 21.2 | 21.9 | 20.9 | 21 | 22.2 | 22.3938 | 22.6 | 18.7 | 22 | 23.7 | 21.9 | 20 | 22.209 | 20.5795 | 20.57 | 27.036 | 45
| 1 | 6 | 22.0812 | 22 | 23.3 | 24 | 21.9 | 23.7 | 21.5 | 20.6 | 19.7 | 21 | 22.6 | 22.9836 | 21.5 | 19.1 | 22 | 24 | 20.8 | 19 | 22.659 | 20.8377 | 22.26 | 22.438
| Mean |  | 22.0358 | 21.8333 | 23.5167 | 24.3333 | 22.0333 | 23.0333 | 20.7833 | 21.6333 | 20.8 | 21.3333 | 22.1667 | 22.2989 | 21.85 | 19.0667 | 22 | 23.7667 | 21.25 | 19.8333 | 21.5007 | 21.6826 | 21.6383 | 23.4565
| Median |  | 22.0178 | 22 | 23.3 | 24 | 22.05 | 22.8 | 20.9 | 21.9 | 21 | 21 | 22.1 | 22.277 | 22 | 19.05 | 22 | 23.8 | 21.25 | 20 | 21.595 | 21.6302 | 21.655 | 22.712
| Std Dev. |  | 0.082797 | 0.40825 | 0.95586 | 0.5164 | 0.33267 | 0.67132 | 0.57417 | 0.64395 | 0.91214 | 0.5164 | 0.41312 | 0.57882 | 0.86197 | 0.30111 | 1.2649 | 0.53541 | 0.45935 | 0.40825 | 0.9438 | 0.92667 | 0.6475 | 2.089
| Rel.Std.Dev. |  | 0.3757% | 1.87% | 4.065% | 2.122% | 1.51% | 2.915% | 2.763% | 2.977% | 4.385% | 2.421% | 1.864% | 2.596% | 3.945% | 1.579% | 5.75% | 2.253% | 2.162% | 2.058% | 4.39% | 4.274% | 2.992% | 8.906%
| PDM3 |  | 0.7188% | -0.2066% | 7.487% | 11.22% | 0.7076% | 5.278% | -5.006% | -1.121% | -4.93% | -2.492% | 1.317% | 1.922% | -0.1304% | -12.85% | 0.5552% | 8.63% | -2.873% | -9.348% | -1.727% | -0.8957% | -1.098% | 7.212%
| Z-Score (Absolute) |  | 0.27 | 0.01 | 2.34 | 3.48 | 0.27 | 1.66 | 1.48 | 0.29 | 1.45 | 0.71 | 0.45 | 0.64 | 0.01 | 3.88 | 0.22 | 2.69 | 0.83 | 2.8 | 0.48 | 0.22 | 0.28 | 2.26
| Analytical results for P in OREAS 605b (Certified Value 624 ppm) | [23 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 1241.6 | 617 | 560 | 1210 | 620 | 610 | 630 | 650 | 640 | 1040 | 607 | 516.1521 | 554 | 596 | 650 | 598 | 636 | < 100 | 697.278 | 610.4926 | < 10 | 1030 | 1
| 1 | 2 | 1241.755 | 616 | 600 | 1210 | 630 | 610 | 640 | 690 | 630 | 1010 | 530 | 510.7432 | 612 | 610 | 600 | 605 | 636 | < 100 | 714.454 | 609.1481 | < 10 | 1090 | #N/A
| 1 | 3 | 1202.555 | 623 | 610 | 1210 | 610 | 620 | 630 | 610 | 640 | 1040 | 546 | 494.9274 | 615 | 591 | 650 | 602 | 644 | < 100 | 720.294 | 610.2877 | < 10 | 1160 | 16
| 1 | 4 | 1268.58 | 610 | 620 | 1210 | 640 | 620 | 620 | 640 | 620 | 1000 | 582 | 497.8794 | 609 | 625 | 650 | 627 | 614 | < 100 | 660.939 | 619.3909 | < 10 | 1200 | 623.5983
| 1 | 5 | 1221.33 | 575 | 610 | 1210 | 630 | 600 | 630 | 650 | 630 | 1000 | 613 | 518.8151 | 621 | 583 | 650 | 631 | 651 | < 100 | 708.618 | 610.6871 | < 10 | 1110 | 46
| 1 | 6 | 1219.18 | 620 | 600 | 1210 | 640 | 620 | 640 | 650 | 610 | 1020 | 569 | 493.3235 | 575 | 607 | 600 | 601 | 652 | < 100 | 733.941 | 616.3425 | < 10 | 1150
| Mean |  | 1232.5 | 610.1667 | 600 | 1210 | 628.3333 | 613.3333 | 631.6667 | 648.3333 | 628.3333 | 1018.3333 | 574.5 | 505.3068 | 597.6667 | 602 | 633.3333 | 610.6667 | 638.8333 |  | 705.9207 | 612.7248 |  | 1123.3333
| Median |  | 1231.465 | 616.5 | 605 | 1210 | 630 | 615 | 630 | 650 | 630 | 1015 | 575.5 | 504.3113 | 610.5 | 601.5 | 650 | 603.5 | 640 |  | 711.536 | 610.5898 |  | 1130
| Std Dev. |  | 23.1047 | 17.7698 | 20.9762 | 0 | 11.6905 | 8.165 | 7.5277 | 25.6255 | 11.6905 | 18.3485 | 32.9166 | 11.2796 | 26.8303 | 15.0732 | 25.8199 | 14.4314 | 14.006 |  | 25.1788 | 4.1326 |  | 59.8888
| Rel.Std.Dev. |  | 1.875% | 2.912% | 3.496% | 0% | 1.861% | 1.331% | 1.192% | 3.953% | 1.861% | 1.802% | 5.73% | 2.232% | 4.489% | 2.504% | 4.077% | 2.363% | 2.192% |  | 3.567% | 0.6745% |  | 5.331%
| PDM3 |  | 97.64% | -2.154% | -3.784% | 94.04% | 0.7593% | -1.646% | 1.294% | 3.966% | 0.7593% | 63.3% | -7.873% | -18.97% | -4.158% | -3.463% | 1.561% | -2.074% | 2.443% |  | 13.2% | -1.744% |  | 80.14%
| Z-Score (Absolute) |  | 18.75 | 0.33 | 0.64 | 18.06 | 0.23 | 0.23 | 0.33 | 0.84 | 0.23 | 12.18 | 1.42 | 3.54 | 0.71 | 0.58 | 0.38 | 0.31 | 0.55 | 17.5 | 2.61 | 0.25 | 18.88 | 15.4
| Analytical results for Pb in OREAS 605b (Certified Value 1510 ppm) | [23 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
0
| 1 | 1 | 1310.1 | 1484 | 1506.85 | 1540 | 1525 | 1500 | 1545 | 1570 | 1550 | 1460 | 1467.3 | 1545.7098 | 1392 | 1340 | 1480 | 1520 | 1140 | 1190 | 1539.5751 | 1463.5804 | 1537 | 1303.416 | 1
| 1 | 2 | 1334.7 | 1457 | 1485.84 | 1540 | 1545 | 1500 | 1535 | 1590 | 1530 | 1500 | 1500.7 | 1504.5403 | 1411 | 1530 | 1500 | 1503 | 1280 | 1018 | 1512.2028 | 1490.0862 | 1515 | 1273.804 | 35
| 1 | 3 | 1243.2 | 1509 | 1491.04 | 1540 | 1530 | 1510 | 1535 | 1570 | 1560 | 1530 | 1356.6 | 1490.9715 | 1430 | 1320 | 1490 | 1501 | 1200 | 1128 | 1521.6673 | 1471.647 | 1502 | 1323.62 | 16
| 1 | 4 | 1387.1 | 1504 | 1507.15 | 1550 | 1530 | 1510 | 1525 | 1560 | 1520 | 1490 | 1332.9 | 1537.8549 | 1448 | 1330 | 1520 | 1502 | 1160 | 1016 | 1442.253 | 1463.479 | 1517 | 1352.899 | 1509.614
| 1 | 5 | 1340.9 | 1411 | 1501.43 | 1550 | 1520 | 1520 | 1530 | 1480 | 1520 | 1530 | 1268.2 | 1530.6198 | 1449 | 1430 | 1490 | 1510 | 1300 | 1032 | 1516.6476 | 1481.1167 | 1490 | 1333.472 | 47
| 1 | 6 | 1349 | 1439 | 1432.87 | 1550 | 1560 | 1570 | 1560 | 1570 | 1510 | 1510 | 1238.1 | 1504.5211 | 1400 | 1530 | 1560 | 1539 | 1220 | 1163 | 1523.9099 | 1467.5193 | 1515 | 1345.703
| Mean |  | 1327.5 | 1467.3333 | 1487.53 | 1545 | 1535 | 1518.3333 | 1538.3333 | 1556.6667 | 1531.6667 | 1503.3333 | 1360.6333 | 1519.0362 | 1421.6667 | 1413.3333 | 1506.6667 | 1512.5 | 1216.6667 | 1091.1667 | 1509.376 | 1472.9048 | 1512.6667 | 1322.1523
| Median |  | 1337.8 | 1470.5 | 1496.235 | 1545 | 1530 | 1510 | 1535 | 1570 | 1525 | 1505 | 1344.75 | 1517.58 | 1420.5 | 1385 | 1495 | 1506.5 | 1210 | 1080 | 1519.1575 | 1469.5831 | 1515 | 1328.546
| Std Dev. |  | 48.2806 | 38.5781 | 28.1223 | 5.4772 | 14.8324 | 26.3944 | 12.5167 | 38.8158 | 19.4079 | 26.5832 | 105.2047 | 21.9465 | 24.3858 | 98.5224 | 29.4392 | 14.8155 | 63.7704 | 78.4714 | 34.178 | 10.6678 | 15.8072 | 29.3901
| Rel.Std.Dev. |  | 3.637% | 2.629% | 1.891% | 0.3545% | 0.9663% | 1.738% | 0.8137% | 2.494% | 1.267% | 1.768% | 7.732% | 1.445% | 1.715% | 6.971% | 1.954% | 0.9795% | 5.241% | 7.192% | 2.264% | 0.7243% | 1.045% | 2.223%
| PDM3 |  | -12.06% | -2.801% | -1.463% | 2.344% | 1.682% | 0.5776% | 1.902% | 3.117% | 1.461% | -0.416% | -9.869% | 0.6241% | -5.826% | -6.378% | -0.1952% | 0.1912% | -19.41% | -27.72% | -0.01577% | -2.432% | 0.2022% | -12.42%
| Z-Score (Absolute) |  | 3.66 | 0.78 | 0.36 | 0.82 | 0.62 | 0.27 | 0.69 | 1.06 | 0.55 | 0.03 | 2.98 | 0.29 | 1.72 | 1.89 | 0.03 | 0.15 | 5.94 | 8.53 | 0.09 | 0.66 | 0.16 | 3.77
| Analytical results for Pd in OREAS 605b (Indicative Value < 50 ppb) | [23 empty] | Indicative
Pd | Round | Replicate | 01: Lab | 1
| No | No | 01: 23 | ppb
|  |  | 01: 5A*MS | 1
1
| 1 | 1 | 01: < 50 | 1
| 1 | 2 | 01: < 50 | 1
| 1 | 3 | 01: < 50 | 16
| 1 | 4 | 01: < 50 | < 50
| 1 | 5 | 01: < 50 | 12
| 1 | 6 | 01: < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for Pr in OREAS 605b (Certified Value 3.01 ppm) | [23 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 23 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 5A*MS | 2
2
| 1 | 1 | 01: 3.3047; 02: 1.8; 03: 3.66; 04: 1.7; 05: 3.22; 06: 2.974 | 1
| 1 | 2 | 01: 3.4063; 02: 1.4; 03: 3.77; 04: 1.6; 05: 2.96; 06: 2.997 | 15
| 1 | 3 | 01: 3.599; 02: 1.3; 03: 3.74; 04: 1.8; 05: 3.08; 06: 3.113 | 16
| 1 | 4 | 01: 3.3652; 02: 1.4; 03: 3.8; 04: 1.8; 05: 2.94; 06: 3.038 | 3.005
| 1 | 5 | 01: 3.4201; 02: 1.5; 03: 3.83; 04: 1.9; 05: 2.9; 06: 3.114 | 48
| 1 | 6 | 01: 3.346; 02: 1.3; 03: 3.83; 04: 1.8; 05: 3.14; 06: 3.003
| Mean |  | 01: 3.4069; 02: 1.45; 03: 3.7717; 04: 1.7667; 05: 3.04; 06: 3.0398
| Median |  | 01: 3.3858; 02: 1.4; 03: 3.785; 04: 1.8; 05: 3.02; 06: 3.0205
| Std Dev. |  | 01: 0.1029; 02: 0.18708; 03: 0.064936; 04: 0.10328; 05: 0.12649; 06: 0.060635
| Rel.Std.Dev. |  | 01: 3.02%; 02: 12.9%; 03: 1.722%; 04: 5.846%; 05: 4.161%; 06: 1.995%
| PDM3 |  | 01: 13.37%; 02: -51.75%; 03: 25.51%; 04: -41.21%; 05: 1.164%; 06: 1.159%
| Z-Score (Absolute) |  | 01: 0.45; 02: 1.95; 03: 0.9; 04: 1.56; 05: 0; 06: 0
| Analytical results for Pt in OREAS 605b (Indicative Value < 10 ppb) | [23 empty] | Indicative
Pt | Round | Replicate | 01: Lab | 1
| No | No | 01: 23 | ppb
|  |  | 01: 5A*MS | 1
1
| 1 | 1 | 01: < 10 | 1
| 1 | 2 | 01: < 10 | 1
| 1 | 3 | 01: < 10 | 16
| 1 | 4 | 01: < 10 | < 10
| 1 | 5 | 01: < 10 | 13
| 1 | 6 | 01: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for Rb in OREAS 605b (Certified Value 39.5 ppm) | [23 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | [5 empty] | 1
2
| 1 | 1 | 43.4482 | 36.9 | 41.9 | 41.1 | 37.6 | 39.6 | 39.8 | 40.6 | 34.4 | 40.07 | 38.4386 | 42.7 | 40.9 | 36.2 | 41.1 | 34.8 | 39.4898 | [5 empty] | 1
| 1 | 2 | 43.2449 | 37.5 | 41.2 | 40.4 | 36.5 | 38.9 | 41.1 | 39.5 | 34.1 | 39.99 | 38.6387 | 46.4 | 41.5 | 32.6 | 39.7 | 35.4 | 39.2989 | [5 empty] | 37
| 1 | 3 | 43.1463 | 35.8 | 43.4 | 40.3 | 38 | 39.8 | 38.3 | 39 | 34.8 | 40.4 | 38.5024 | 38.4 | 41.3 | 36.4 | 39.9 | 34.6 | 40.0005 | [5 empty] | 16
| 1 | 4 | 43.1781 | 36.7 | 40.7 | 39.2 | 40.1 | 40.1 | 41 | 38.6 | 33.4 | 41.08 | 38.3917 | 44.6 | 40.8 | 35.8 | 38.8 | 35.1 | 39.1598 | [5 empty] | 39.5401
| 1 | 5 | 43.6439 | 35.8 | 40.8 | 37.1 | 40.1 | 37.8 | 40.6 | 38.4 | 33.8 | 40.76 | 38.7031 | 45.9 | 39.1 | 37.2 | 40.2 | 35.8 | 40.9342 | [5 empty] | 49
| 1 | 6 | 43.0678 | 38.3 | 39.8 | 37.7 | 42.7 | 37.7 | 39.6 | 36.9 | 33.1 | 40.72 | 38.4923 | 42 | 41.3 | 37.2 | 40.7 | 35.5 | 39.6486
| Mean |  | 43.2882 | 36.8333 | 41.3 | 39.3 | 39.1667 | 38.9833 | 40.0667 | 38.8333 | 33.9333 | 40.5033 | 38.5278 | 43.3333 | 40.8167 | 35.9 | 40.0667 | 35.2 | 39.7553
| Median |  | 43.2115 | 36.8 | 41 | 39.75 | 39.05 | 39.25 | 40.2 | 38.8 | 33.95 | 40.56 | 38.4974 | 43.65 | 41.1 | 36.3 | 40.05 | 35.25 | 39.5692
| Std Dev. |  | 0.21674 | 0.97502 | 1.2361 | 1.6037 | 2.2429 | 1.0342 | 1.0614 | 1.2307 | 0.6314 | 0.42599 | 0.11952 | 2.9703 | 0.88185 | 1.71 | 0.80664 | 0.45166 | 0.64726
| Rel.Std.Dev. |  | 0.5007% | 2.647% | 2.993% | 4.081% | 5.727% | 2.653% | 2.649% | 3.169% | 1.861% | 1.052% | 0.3102% | 6.855% | 2.161% | 4.763% | 2.013% | 1.283% | 1.628%
| PDM3 |  | 9.479% | -6.846% | 4.451% | -0.6073% | -0.9445% | -1.408% | 1.332% | -1.788% | -14.18% | 2.436% | -2.56% | 9.593% | 3.228% | -9.206% | 1.332% | -10.98% | 0.5442%
| Z-Score (Absolute) |  | 2.23 | 1.38 | 1.12 | 0 | 0.07 | 0.18 | 0.43 | 0.26 | 3.01 | 0.67 | 0.43 | 2.26 | 0.85 | 1.91 | 0.43 | 2.3 | 0.26
| Analytical results for Re in OREAS 605b (Indicative Value 11.4 ppb) | [23 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 23 | [8 empty] | ppb
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 5A*MS | [8 empty] | 1
1
| 1 | 1 | < 50 | 3 | 9 | 11 | 10 | 10 | 12 | < 50 | 12 | < 50 | 15 | 9 | < 100 | 14 | [8 empty] | 1
| 1 | 2 | < 50 | 11 | 11 | 15 | 7 | 11 | 13 | < 50 | 8 | < 50 | 10 | 11 | < 100 | 13 | [8 empty] | 8
| 1 | 3 | < 50 | 13 | 9 | 23 | 8 | 10 | 13 | < 50 | 13 | < 50 | 11 | 9 | < 100 | 16 | [8 empty] | 16
| 1 | 4 | < 50 | 18 | 10 | 21 | 11 | 10 | 9 | < 50 | 9 | < 50 | 9 | 8 | < 100 | 15 | [8 empty] | 11.3704
| 1 | 5 | < 50 | 16 | 9 | 25 | 11 | 13 | 12 | < 50 | 13 | < 50 | 13 | 9 | < 100 | 15 | [8 empty] | 14
| 1 | 6 | < 50 | 11 | 10 | 9 | 10 | 16 | 12 | < 50 | 10 | < 50 | 13 | 9 | < 100 | 13
| Mean |  |  | 12 | 9.6667 | 17.3333 | 9.5 | 11.6667 | 11.8333 |  | 10.8333 |  | 11.8333 | 9.1667 |  | 14.3333
| Median |  |  | 12 | 9.5 | 18 | 10 | 10.5 | 12 |  | 11 |  | 12 | 9 |  | 14.5
| Std Dev. |  |  | 5.2154 | 0.8165 | 6.6232 | 1.6432 | 2.4221 | 1.472 |  | 2.137 |  | 2.2286 | 0.98319 |  | 1.2111
| Rel.Std.Dev. |  |  | 43.46% | 8.447% | 38.21% | 17.3% | 20.76% | 12.44% |  | 19.73% |  | 18.83% | 10.73% |  | 8.449%
| PDM3 |  |  | 5.537% | -14.98% | 52.44% | -16.45% | 2.606% | 4.072% |  | -4.723% |  | 4.072% | -19.38% |  | 26.06%
| Z-Score (Absolute) |  | 3.65 | 0.02 | 0.63 | 1.51 | 0.67 | 0.07 | 0.02 | 3.65 | 0.3 | 3.65 | 0.02 | 0.77 | 10.63 | 0.67
| Analytical results for Rh in OREAS 605b (Indicative Value 40 ppb) | [23 empty] | Indicative
Rh | Round | Replicate | 02: Lab | [8 empty] | 1
| No | No | 02: 23 | [8 empty] | ppb
|  |  | 02: 5A*MS | [8 empty] | 1
1
| 1 | 1 | 02: 40 | [8 empty] | 1
| 1 | 2 | 02: 40 | [8 empty] | 5
| 1 | 3 | 02: 40 | [8 empty] | 16
| 1 | 4 | 02: 40 | [8 empty] | 40
| 1 | 5 | 02: 40 | [8 empty] | 11
| 1 | 6 | 02: 40
| Mean |  | 02: 40
| Median |  | 02: 40
| Std Dev. |  | 02: 0
Rel.Std.Dev. | bar: 02
PDM3 | bar: 02
Z-Score (Absolute) | bar: 02
| Analytical results for Ru in OREAS 605b (Indicative Value < 10 ppb) | [23 empty] | Indicative
Ru | Round | Replicate | 02: Lab | [8 empty] | 1
| No | No | 02: 23 | [8 empty] | ppb
|  |  | 02: 5A*MS | [8 empty] | 1
1
| 1 | 1 | 02: < 10 | [8 empty] | 1
| 1 | 2 | 02: < 10 | [8 empty] | 6
| 1 | 3 | 02: < 10 | [8 empty] | 16
| 1 | 4 | 02: < 10 | [8 empty] | < 10
| 1 | 5 | 02: < 10 | [8 empty] | 12
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for S in OREAS 605b (Certified Value 9.96 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | PF*OES | 5A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 9.4094 | 9.43 | 10.19 |  |  | 10 | > 10 | 9.71 | 10.5 | 9.01 | 10.4556 | > 10 | 10.5 | 9.95 | > 5 | > 5 | 9.85 | 9.9262 | 10.5867 | 8.804 | > 10 |  | 1
| 1 | 2 | 9.4301 | 9.63 | 10.18 |  |  | 9.97 | > 10 | 9.58 | 10.1 | 8.92 | 10.067 | > 10 | 10.4 | 10.1 | > 5 | > 5 | 9.78 | 9.9234 | 10.7258 | 8.771 | > 10 |  | #N/A
| 1 | 3 | 9.3728 | 9.68 | 10.18 |  |  | 9.91 | 9.78 | 9.75 | 10.1 | 8.92 | 10.2412 | > 10 | 9.94 | 9.93 | > 5 | > 5 | 9.99 | 9.8299 | 10.8509 | 8.75 | 9.961 |  | 16
| 1 | 4 | 9.141 | 9.76 | 10.2 | 10 |  | 9.86 | > 10 | 9.35 | 10.1 | 9.06 | 10.1199 | > 10 | 10.3 | 10.1 | > 5 | > 5 | 8.87 | 9.7672 | 10.7973 | 8.743 | 9.965 |  | 9.9623
| 1 | 5 | 8.6803 | 9.79 | 10.19 | 9.88 |  | 9.88 | > 10 | 9.55 | 10.2 | 8.82 | 10.3709 | > 10 | 9.77 | 10.1 | > 5 | > 5 | 9.78 | 9.8011 | 10.8914 | 8.743 | > 10 |  | 50
| 1 | 6 | 9.3244 | 9.56 | 10.2 | 9.89 | 10.1 | > 10 | > 10 | 9.41 | 10.1 | 8.91 | 10.2023 | > 10 | 10.3 | 10.1 | > 5 | > 5 | 9.19 | 9.9607 | 10.8497 | 8.767 | > 10
| Mean |  | 9.2263 | 9.6417 | 10.19 | 9.9233 | 10.1 | 9.924 | 9.78 | 9.5583 | 10.1833 | 8.94 | 10.2428 |  | 10.2017 | 10.0467 |  |  | 9.5767 | 9.8681 | 10.7836 | 8.763 | 9.963
| Median |  | 9.3486 | 9.655 | 10.19 | 9.89 | 10.1 | 9.91 | 9.78 | 9.565 | 10.1 | 8.92 | 10.2217 |  | 10.3 | 10.1 |  |  | 9.78 | 9.8767 | 10.8235 | 8.7585 | 9.963
| Std Dev. |  | 0.28689 | 0.13348 | 0.0089443 | 0.066583 |  | 0.059414 |  | 0.15855 | 0.16021 | 0.084143 | 0.14791 |  | 0.28372 | 0.082865 |  |  | 0.44207 | 0.078929 | 0.1121 | 0.023367 | 0.0028284
| Rel.Std.Dev. |  | 3.11% | 1.384% | 0.08777% | 0.671% |  | 0.5987% |  | 1.659% | 1.573% | 0.9412% | 1.444% |  | 2.781% | 0.8248% |  |  | 4.616% | 0.7998% | 1.04% | 0.2667% | 0.02839%
| PDM3 |  | -7.387% | -3.218% | 2.286% | -0.3911% | 1.382% | -0.3844% | -1.83% | -4.055% | 2.219% | -10.26% | 2.816% |  | 2.403% | 0.8469% |  |  | -3.871% | -0.9457% | 8.244% | -12.04% | 0.007058%
| Z-Score (Absolute) |  | 1.73 | 0.7 | 0.66 | 0 | 0.44 | 0 | 0.36 | 0.91 | 0.64 | 2.44 | 0.79 | NA | 0.69 | 0.3 | NA | NA | 0.86 | 0.14 | 2.13 | 2.88 | 0.1
| Analytical results for Sb in OREAS 605b (Certified Value 405 ppm) | [23 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 331.64 | 387.3 | 456.16 | 430 | 403 | 385 | 422 | 427 | 429 | 397 | 406.59 | 395.4039 | 386 | 477 | 448.2 | 389 | 408 | 395.9726 | 389.1655 | 403.1 | 363.565 |  | 1
| 1 | 2 | 334.36 | 383.9 | 439.53 | 430 | 417 | 394 | 419 | 432 | 415 | 402 | 399.23 | 387.7316 | 388.29 | 478 | 428.3 | 398 | 393 | 395.1797 | 375.1232 | 400.8 | 341.6 |  | 39
| 1 | 3 | 333.69 | 377.9 | 448.09 | 430 | 410 | 396 | 412 | 425 | 426 | 401 | 396.85 | 374.4524 | 393.75 | 479 | 434.1 | 393 | 406 | 403.6435 | 352.2327 | 405 | 374.874 |  | 16
| 1 | 4 | 340.75 | 364 | 449.64 | 440 | 413 | 387 | 415 | 454 | 419 | 393 | 407.27 | 381.2094 | 385.87 | 473 | 425.4 | 393 | 390 | 396.4811 | 348.4277 | 409.1 | 350.763 |  | 404.6231
| 1 | 5 | 336.29 | 366 | 438.25 | 430 | 433 | 392 | 410 | 455 | 417 | 398 | 409.25 | 382.7868 | 389.9 | 466 | 435.6 | 402 | 407 | 410.9694 | 386.861 | 402 | 344.559 |  | 51
| 1 | 6 | 338.07 | 385.4 | 447.63 | 430 | 445 | 414 | 420 | 448 | 418 | 402 | 409.1 | 394.1314 | 383.9 | 479 | 444.3 | 393 | 390 | 397.4925 | 416.4812 | 405.6 | 354.723
| Mean |  | 335.8 | 377.4167 | 446.55 | 431.6667 | 420.1667 | 394.6667 | 416.3333 | 440.1667 | 420.6667 | 398.8333 | 404.715 | 385.9526 | 387.9517 | 475.3333 | 435.9833 | 394.6667 | 399 | 399.9565 | 378.0485 | 404.2667 | 355.014
| Median |  | 335.325 | 380.9 | 447.86 | 430 | 415 | 393 | 417 | 440 | 418.5 | 399.5 | 406.93 | 385.2592 | 387.145 | 477.5 | 434.85 | 393 | 399.5 | 396.9868 | 380.9921 | 404.05 | 352.743
| Std Dev. |  | 3.278 | 10.1393 | 6.6874 | 4.0825 | 15.7533 | 10.3473 | 4.761 | 13.732 | 5.5377 | 3.5449 | 5.3251 | 8.0518 | 3.5239 | 5.0859 | 8.8673 | 4.5898 | 8.8544 | 6.1922 | 25.4194 | 2.973 | 12.4494
| Rel.Std.Dev. |  | 0.9762% | 2.687% | 1.498% | 0.9457% | 3.749% | 2.622% | 1.144% | 3.12% | 1.316% | 0.8888% | 1.316% | 2.086% | 0.9083% | 1.07% | 2.034% | 1.163% | 2.219% | 1.548% | 6.724% | 0.7354% | 3.507%
| PDM3 |  | -17.01% | -6.724% | 10.36% | 6.684% | 3.841% | -2.461% | 2.894% | 8.784% | 3.965% | -1.431% | 0.02271% | -4.614% | -4.12% | 17.48% | 7.75% | -2.461% | -1.39% | -1.153% | -6.568% | -0.0881% | -12.26%
| Z-Score (Absolute) |  | 2.14 | 0.75 | 1.55 | 1.06 | 0.67 | 0.18 | 0.55 | 1.34 | 0.69 | 0.04 | 0.16 | 0.47 | 0.4 | 2.51 | 1.2 | 0.18 | 0.03 | 0 | 0.73 | 0.14 | 1.5
| Analytical results for Sc in OREAS 605b (Certified Value 6.31 ppm) | [23 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 26 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 7.3245 | 6 | 5.3 | < 5 | 6.5 | 6.1 | 6.4 | 6.5 | 6.5 | 7 | 6.9 | 5.8648 | 6 | 6.7 | 6 | 6.9 | 5.31 | 6 | 6.4812 | 6.2994 | 5.751 |  | 1
| 1 | 2 | 7.1305 | 6 | 5.4 | < 5 | 6.6 | 6.7 | 6.4 | 6.6 | 6.5 | 7 | 6.6 | 6.0327 | 7 | 6.8 | 6 | 6.9 | 5.16 | 5.8 | 6.3536 | 6.1041 | 5.539 |  | 40
| 1 | 3 | 7.5645 | 6 | 6.1 | < 5 | 6.6 | 7.2 | 6.3 | 5.8 | 6.3 | 7 | 6.4 | 6.2653 | 7 | 6.8 | 7 | 6.7 | 5.29 | 5.6 | 6.4746 | 6.4193 | 5.876 |  | 16
| 1 | 4 | 7.4316 | 6 | 5.8 | < 5 | 6.5 | 5.8 | 6.7 | 6.8 | 6 | 7 | 7.1 | 5.9164 | 7 | 6.9 | 7 | 7.2 | 5.31 | 6 | 6.2018 | 6.3555 | 5.792 |  | 6.3135
| 1 | 5 | 7.7083 | 6 | 5.4 | < 5 | 6.5 | 5.7 | 6.3 | 6.7 | 6.3 | 7 | 7 | 6.3794 | 7 | 6.6 | 7 | 7.2 | 5.52 | 6.2 | 6.435 | 6.0245 | 5.55 |  | 52
| 1 | 6 | 7.5219 | 6 | 5.2 | < 5 | 6.6 | 6.6 | 6.3 | 6.2 | 5.9 | 7 | 7 | 6.0725 | 7 | 6.9 | 7 | 6.7 | 5.23 | 5.8 | 6.4262 | 6.1374 | 5.171
| Mean |  | 7.4469 | 6 | 5.5333 |  | 6.55 | 6.35 | 6.4 | 6.4333 | 6.25 | 7 | 6.8333 | 6.0885 | 6.8333 | 6.7833 | 6.6667 | 6.9333 | 5.3033 | 5.9 | 6.3954 | 6.2234 | 5.6132
| Median |  | 7.4768 | 6 | 5.4 |  | 6.55 | 6.35 | 6.35 | 6.55 | 6.3 | 7 | 6.95 | 6.0526 | 7 | 6.8 | 7 | 6.9 | 5.3 | 5.9 | 6.4306 | 6.2184 | 5.6505
| Std Dev. |  | 0.20152 | 0 | 0.34448 |  | 0.054772 | 0.58224 | 0.15492 | 0.37238 | 0.251 | 0 | 0.27325 | 0.19946 | 0.40825 | 0.1169 | 0.5164 | 0.22509 | 0.12094 | 0.20976 | 0.10524 | 0.1567 | 0.25498
| Rel.Std.Dev. |  | 2.706% | 0% | 6.226% |  | 0.8362% | 9.169% | 2.421% | 5.788% | 4.016% | 0% | 3.999% | 3.276% | 5.974% | 1.723% | 7.746% | 3.247% | 2.28% | 3.555% | 1.646% | 2.518% | 4.542%
| PDM3 |  | 17.95% | -4.966% | -12.36% |  | 3.746% | 0.578% | 1.37% | 1.898% | -1.006% | 10.87% | 8.234% | -3.564% | 8.234% | 7.442% | 5.594% | 9.818% | -16% | -6.55% | 1.297% | -1.428% | -11.09%
| Z-Score (Absolute) |  | 1.71 | NA | 1.27 | 5.99 | 0.31 | 0 | 0.08 | 0.13 | 0.16 | NA | 0.75 | 0.41 | NA | 0.67 | NA | 0.91 | 1.63 | 0.7 | 0.07 | 0.2 | 1.15
Results from laboratories 2, 10, 14 & 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 605b (Certified Value 87 ppm) | [23 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 80 | 87.5 | 93 | 90 | 77 | 93 | 94 | 89 | 85.7 | 82.8255 | 86.3 | 98.8 | 75 | 92 | 76.4 | 87.0476 | 88.066 | [5 empty] | 1
| 1 | 2 | 83 | 86.7 | 92 | 92 | 75 | 97 | 92 | 90 | 84.4 | 83.5111 | 84.1 | 96.9 | 65 | 90 | 74.7 | 85.3794 | 82.6865 | [5 empty] | 41
| 1 | 3 | 85 | 90.8 | 94 | 73 | 79 | 88 | 93 | 91 | 86.7 | 80.1537 | 77.9 | 98.8 | 75 | 86 | 73.6 | 91.6785 | 89.2173 | [5 empty] | 16
| 1 | 4 | 79 | 90.8 | 96 | 88 | 79 | 98 | 87 | 91 | 84.6 | 81.7899 | 96.2 | 96.9 | 75 | 89 | 74 | 89.6152 | 87.0174 | [5 empty] | 86.7136
| 1 | 5 | 74 | 89.3 | 95 | 94 | 82 | 99 | 91 | 89 | 84.3 | 81.3633 | 86.3 | 95.3 | 80 | 95 | 74.8 | 91.8505 | 91.7861 | [5 empty] | 53
| 1 | 6 | 81 | 87.3 | 96 | 92 | 86 | 96 | 89 | 90 | 84.5 | 81.531 | 67.6 | 97.7 | 80 | 91 | 73.3 | 89.2201 | 84.6469
| Mean |  | 80.3333 | 88.7333 | 94.3333 | 88.1667 | 79.6667 | 95.1667 | 91 | 90 | 85.0333 | 81.8624 | 83.0667 | 97.4 | 75 | 90.5 | 74.4667 | 89.1319 | 87.2367
| Median |  | 80.5 | 88.4 | 94.5 | 91 | 79 | 96.5 | 91.5 | 90 | 84.55 | 81.6604 | 85.2 | 97.3 | 75 | 90.5 | 74.35 | 89.4176 | 87.5417
| Std Dev. |  | 3.7771 | 1.8206 | 1.633 | 7.705 | 3.8816 | 4.0702 | 2.6077 | 0.89443 | 0.96264 | 1.1773 | 9.5956 | 1.3357 | 5.4772 | 3.0166 | 1.1165 | 2.5512 | 3.2486
| Rel.Std.Dev. |  | 4.702% | 2.052% | 1.731% | 8.739% | 4.872% | 4.277% | 2.866% | 0.9938% | 1.132% | 1.438% | 11.55% | 1.371% | 7.303% | 3.333% | 1.499% | 2.862% | 3.724%
| PDM3 |  | -7.358% | 2.329% | 8.787% | 1.676% | -8.127% | 9.748% | 4.943% | 3.79% | -1.938% | -5.594% | -4.206% | 12.32% | -13.51% | 4.367% | -14.12% | 2.789% | 0.6033%
| Z-Score (Absolute) |  | 1.04 | 0.07 | 0.82 | 0 | 1.12 | 0.93 | 0.37 | 0.24 | 0.41 | 0.83 | 0.67 | 1.22 | 1.74 | 0.31 | 1.81 | 0.13 | 0.12
| Analytical results for Sm in OREAS 605b (Certified Value 2.15 ppm) | [23 empty] | Certified
Sm | Round | Replicate | 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | [5 empty] | 1
| No | No | 02: 01; 03: 03; 05: 11; 06: 14; 07: 16; 08: 23 | [5 empty] | ppm
|  |  | 02: 4A*MS; 03: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 5A*MS | [5 empty] | 2
2
| 1 | 1 | 02: 2.5598; 03: 2; 05: 2.26; 06: 2.1; 07: 2.1; 08: 1.973 | [5 empty] | 1
| 1 | 2 | 02: 2.4181; 03: 2; 05: 2.25; 06: 2.1; 07: 2; 08: 2.009 | [5 empty] | 12
| 1 | 3 | 02: 2.4979; 03: 1.8; 05: 2.14; 06: 2.1; 07: 2.25; 08: 1.972 | [5 empty] | 16
| 1 | 4 | 02: 2.5396; 03: 1.9; 05: 2.39; 06: 2.2; 07: 2.15; 08: 1.939 | [5 empty] | 2.1457
| 1 | 5 | 02: 2.6542; 03: 1.7; 05: 2.27; 06: 2.1; 07: 2.2; 08: 1.942 | [5 empty] | 54
| 1 | 6 | 02: 2.5249; 03: 1.5; 05: 2.2; 06: 2.3; 07: 2.25; 08: 1.956
| Mean |  | 02: 2.5324; 03: 1.8167; 05: 2.2517; 06: 2.15; 07: 2.1583; 08: 1.9652
| Median |  | 02: 2.5322; 03: 1.85; 05: 2.255; 06: 2.1; 07: 2.175; 08: 1.964
| Std Dev. |  | 02: 0.077394; 03: 0.19408; 05: 0.083287; 06: 0.083666; 07: 0.09704; 08: 0.025826
| Rel.Std.Dev. |  | 02: 3.056%; 03: 10.68%; 05: 3.699%; 06: 3.891%; 07: 4.496%; 08: 1.314%
| PDM3 |  | 02: 18.02%; 03: -15.33%; 05: 4.938%; 06: 0.2001%; 07: 0.5885%; 08: -8.414%
| Z-Score (Absolute) |  | 02: 1.78; 03: 1.59; 05: 0.46; 06: 0.02; 07: 0.02; 08: 0.89
| Analytical results for Sn in OREAS 605b (Certified Value 6.79 ppm) | [23 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 6.4 | 7.7 | < 10 | 6.4 | 6.6 | 6.3 | 7.2 | 7 | 6.4 | 6.9 | 6.9391 | 5.6 | 6.8 | 8 | 7.9 | 6.57 | < 10 | 6.6733 | < 20 | < 5 | < 10 |  | 1
| 1 | 2 | 6.4 | 6.6 | < 10 | 6.5 | 6.6 | 6.3 | 7.7 | 7 | 6.2 | 6.8 | 6.8932 | 6 | 6.4 | 7 | 7.4 | 6.66 | < 10 | 6.4303 | < 20 | < 5 | < 10 |  | 26
| 1 | 3 | 6.5 | 6.7 | < 10 | 6.5 | 7.1 | 6.3 | 6.8 | 7.1 | 6.2 | 7 | 6.9841 | 5.5 | 6.6 | 8 | 7.4 | 6.71 | < 10 | 6.6172 | < 20 | < 5 | < 10 |  | 16
| 1 | 4 | 6.5 | 7.2 | < 10 | 6.7 | 6.7 | 6.6 | 7.5 | 6.7 | 6.4 | 6.9 | 6.7696 | 5.5 | 6.3 | 8 | 7.3 | 6.63 | < 10 | 6.4778 | < 20 | < 5 | < 10 |  | 6.7852
| 1 | 5 | 6.3 | 7.2 | < 10 | 6.7 | 6.8 | 6.5 | 7.5 | 7 | 6.3 | 7 | 6.7644 | 5.7 | 6.1 | 8 | 7.4 | 6.71 | < 10 | 6.8958 | < 20 | < 5 | < 10 |  | 55
| 1 | 6 | 6.7 | 7.3 | < 10 | 6.7 | 6.8 | 6.3 | 7.6 | 6.5 | 6.2 | 6.9 | 6.8427 | 5.2 | 6.7 | 8 | 7.6 | 6.52 | < 10 | 6.5858 | < 20 | < 5 | < 10
| Mean |  | 6.4667 | 7.1167 |  | 6.5833 | 6.7667 | 6.3833 | 7.3833 | 6.8833 | 6.2833 | 6.9167 | 6.8655 | 5.5833 | 6.4833 | 7.8333 | 7.5 | 6.6333 |  | 6.6134
| Median |  | 6.45 | 7.2 |  | 6.6 | 6.75 | 6.3 | 7.5 | 7 | 6.25 | 6.9 | 6.8679 | 5.55 | 6.5 | 8 | 7.4 | 6.645 |  | 6.6015
| Std Dev. |  | 0.13663 | 0.40702 |  | 0.13292 | 0.18619 | 0.13292 | 0.33116 | 0.23166 | 0.098319 | 0.075277 | 0.089655 | 0.26394 | 0.26394 | 0.40825 | 0.21909 | 0.076594 |  | 0.16493
| Rel.Std.Dev. |  | 2.113% | 5.719% |  | 2.019% | 2.752% | 2.082% | 4.485% | 3.366% | 1.565% | 1.088% | 1.306% | 4.727% | 4.071% | 5.212% | 2.921% | 1.155% |  | 2.494%
| PDM3 |  | -4.694% | 4.886% |  | -2.975% | -0.2725% | -5.922% | 8.816% | 1.447% | -7.396% | 1.938% | 1.184% | -17.71% | -4.448% | 15.45% | 10.54% | -2.238% |  | -2.532%
| Z-Score (Absolute) |  | 0.28 | 1.1 | 3.4 | 0.03 | 0.36 | 0.46 | 1.67 | 0.61 | 0.67 | 0.68 | 0.57 | 2.16 | 0.24 | NA | 1.92 | 0.07 | 3.4 | 0.03 | 7.24 | 8.73 | 3.4
Results from laboratory 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 605b (Certified Value 402 ppm) | [23 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 394.0704 | 506.4 | 204 | 157 | 369 | 396 | 418 | 359 | 415 | 438 | 584.89 | 489.2493 | 275 | 312 | 355.8 | 278 | 206 | 461.628 | 490.9991 | 525 | 311.051 |  | 1
| 1 | 2 | 398.6719 | 505.5 | 201 | 158 | 378 | 392 | 398 | 376 | 377 | 425 | 584.21 | 490.0406 | 273 | 310 | 353 | 286 | 219 | 459.106 | 499.6928 | 531.4 | 298.962 |  | 27
| 1 | 3 | 393.6591 | 528.1 | 200 | 156 | 353 | 418 | 416 | 313 | 412 | 427 | 541.16 | 488.2316 | 272 | 323 | 342.4 | 308 | 177 | 469.5791 | 504.6987 | 525.8 | 311.028 |  | 16
| 1 | 4 | 394.6258 | 517.3 | 200 | 157 | 363 | 383 | 405 | 415 | 404 | 421 | 535.41 | 479.7225 | 328 | 338 | 329.1 | 293 | 271 | 454.7404 | 484.3635 | 527.9 | 316.151 |  | 401.6515
| 1 | 5 | 400.6854 | 483.2 | 210 | 155 | 385 | 384 | 415 | 431 | 405 | 427 | 540.13 | 488.0987 | 306 | 308 | 348 | 307 | 214 | 487.171 | 479.5052 | 525.6 | 287.361 |  | 56
| 1 | 6 | 402.4983 | 541.4 | 184 | 155 | 395 | 463 | 419 | 399 | 399 | 422 | 535.1 | 478.8462 | 315 | 309 | 338.2 | 285 | 194 | 474.6627 | 494.151 | 528.1 | 296.005
| Mean |  | 397.3685 | 513.65 | 199.8333 | 156.3333 | 373.8333 | 406 | 411.8333 | 382.1667 | 402 | 426.6667 | 553.4833 | 485.6981 | 294.8333 | 316.6667 | 344.4167 | 292.8333 | 213.5 | 467.8145 | 492.235 | 527.3 | 303.4263
| Median |  | 396.6489 | 511.85 | 200.5 | 156.5 | 373.5 | 394 | 415.5 | 387.5 | 404.5 | 426 | 540.645 | 488.1651 | 290.5 | 311 | 345.2 | 289.5 | 210 | 465.6035 | 492.575 | 526.85 | 304.995
| Std Dev. |  | 3.7729 | 20.191 | 8.6352 | 1.2111 | 15.2632 | 30.6659 | 8.4242 | 42.6728 | 13.5351 | 6.0882 | 24.1879 | 5.026 | 24.5879 | 11.7927 | 9.9339 | 12.3194 | 31.9547 | 11.9141 | 9.3826 | 2.3766 | 11.0539
| Rel.Std.Dev. |  | 0.9495% | 3.931% | 4.321% | 0.7747% | 4.083% | 7.553% | 2.046% | 11.17% | 3.367% | 1.427% | 4.37% | 1.035% | 8.34% | 3.724% | 2.884% | 4.207% | 14.97% | 2.547% | 1.906% | 0.4507% | 3.643%
| PDM3 |  | -1.066% | 27.88% | -50.25% | -61.08% | -6.926% | 1.083% | 2.535% | -4.851% | 0.08676% | 6.228% | 37.8% | 20.93% | -26.59% | -21.16% | -14.25% | -27.09% | -46.84% | 16.47% | 22.55% | 31.28% | -24.46%
| Z-Score (Absolute) |  | 0 | 0.89 | 1.51 | 1.84 | 0.18 | 0.07 | 0.11 | 0.12 | 0.04 | 0.22 | 1.19 | 0.67 | 0.78 | 0.62 | 0.4 | 0.8 | 1.4 | 0.54 | 0.72 | 0.99 | 0.72
| Analytical results for Ta in OREAS 605b (Certified Value 0.55 ppm) | [23 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.55393 | 0.55 | 0.6 | < 5 | 0.54 | 0.49 | 0.54 | 0.55 | 0.56 | 0.64 | 0.59 | 0.29635 | 0.6 | 0.44 | 0.6 | 0.82 | 0.892 | 1.516 | < 10 |  |  |  | 1
| 1 | 2 | 0.56493 | 0.53 | 0.5 | < 5 | 0.55 | 0.5 | 0.57 | 0.56 | 0.56 | 0.61 | 0.57 | 0.26823 | 0.5 | 0.48 | 0.5 | 0.83 | 0.778 | 1.512 | < 10 |  |  |  | 13
| 1 | 3 | 0.56375 | 0.55 | 0.6 | < 5 | 0.53 | 0.51 | 0.55 | 0.54 | 0.56 | 0.65 | 0.57 | 0.28857 | 0.6 | 0.43 | 0.5 | 0.76 | 0.858 | 1.582 | < 10 |  |  |  | 16
| 1 | 4 | 0.53187 | 0.53 | 0.6 | < 5 | 0.56 | 0.54 | 0.55 | 0.58 | 0.54 | 0.62 | 0.59 | 0.26226 | 0.6 | 0.43 | 0.4 | 0.76 | 0.96 | 1.533 | < 10 |  |  |  | 0.54771
| 1 | 5 | 0.5275 | 0.52 | 0.6 | < 5 | 0.54 | 0.57 | 0.51 | 0.58 | 0.54 | 0.64 | 0.58 | 0.20236 | 0.6 | 0.49 | 0.6 | 0.78 | 0.899 | 1.569 | < 10 |  |  |  | 57
| 1 | 6 | 0.52883 | 0.5 | 0.5 | < 5 | 0.56 | 0.51 | 0.55 | 0.58 | 0.54 | 0.61 | 0.59 | 0.20441 | 0.6 | 0.48 | 0.5 | 0.8 | 0.876 | 1.56 | < 10
| Mean |  | 0.54513 | 0.53 | 0.56667 |  | 0.54667 | 0.52 | 0.545 | 0.565 | 0.55 | 0.62833 | 0.58167 | 0.2537 | 0.58333 | 0.45833 | 0.51667 | 0.79167 | 0.87717 | 1.5453
| Median |  | 0.5429 | 0.53 | 0.6 |  | 0.545 | 0.51 | 0.55 | 0.57 | 0.55 | 0.63 | 0.585 | 0.26525 | 0.6 | 0.46 | 0.5 | 0.79 | 0.884 | 1.5465
| Std Dev. |  | 0.017711 | 0.018974 | 0.05164 |  | 0.012111 | 0.029665 | 0.019748 | 0.017607 | 0.010954 | 0.017224 | 0.0098319 | 0.04095 | 0.040825 | 0.027869 | 0.075277 | 0.029944 | 0.0596 | 0.029132
| Rel.Std.Dev. |  | 3.249% | 3.58% | 9.113% |  | 2.215% | 5.705% | 3.624% | 3.116% | 1.992% | 2.741% | 1.69% | 16.14% | 6.999% | 6.08% | 14.57% | 3.782% | 6.795% | 1.885%
| PDM3 |  | -0.4708% | -3.234% | 3.46% |  | -0.1911% | -5.06% | -0.4954% | 3.156% | 0.4175% | 14.72% | 6.199% | -53.68% | 6.503% | -16.32% | -5.668% | 44.54% | 60.15% | 182.1%
| Z-Score (Absolute) |  | 0.15 | 0.34 | NA | 24.18 | 0.13 | 0.47 | 0.16 | 0.09 | 0.09 | 0.88 | 0.3 | 3.78 | NA | 1.23 | NA | 2.92 | 3.98 | 12.3 | 55.3
Results from laboratories 3, 14 & 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 605b (Certified Value 0.22 ppm) | [23 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 20; 07: 21; 08: 23 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES/MS; 08: 5A*OES |  |  |  | 2
2
| 1 | 1 | 01: 0.40512; 02: 0.4; 03: 0.25; 04: 0.2; 05: 0.22; 06: 0.27; 07: 0.219; 08: 1.386 |  |  |  | 1
| 1 | 2 | 01: 0.38634; 02: 0.2; 03: 0.24; 04: 0.2; 05: 0.2; 06: 0.26; 07: 0.202; 08: 1.198 |  |  |  | 14
| 1 | 3 | 01: 0.38063; 02: 0.2; 03: 0.25; 04: 0.2; 05: 0.22; 06: 0.25; 07: 0.207; 08: 1.044 |  |  |  | 16
| 1 | 4 | 01: 0.38586; 02: 0.2; 03: 0.25; 04: 0.3; 05: 0.22; 06: 0.25; 07: 0.201; 08: 1.028 |  |  |  | 0.2185
| 1 | 5 | 01: 0.41354; 02: 0.1; 03: 0.25; 04: 0.2; 05: 0.22; 06: 0.26; 07: 0.213; 08: 1.24 |  |  |  | 58
| 1 | 6 | 01: 0.36725; 02: 0.1; 03: 0.24; 04: 0.2; 05: 0.22; 06: 0.28; 07: 0.214; 08: 1.265
| Mean |  | 01: 0.38979; 02: 0.2; 03: 0.24667; 04: 0.21667; 05: 0.21667; 06: 0.26167; 07: 0.20933; 08: 1.1935
| Median |  | 01: 0.3861; 02: 0.2; 03: 0.25; 04: 0.2; 05: 0.22; 06: 0.26; 07: 0.21; 08: 1.219
| Std Dev. |  | 01: 0.016842; 02: 0.10954; 03: 0.005164; 04: 0.040825; 05: 0.008165; 06: 0.01169; 07: 0.007174; 08: 0.13718
| Rel.Std.Dev. |  | 01: 4.321%; 02: 54.77%; 03: 2.094%; 04: 18.84%; 05: 3.768%; 06: 4.468%; 07: 3.427%; 08: 11.49%
| PDM3 |  | 01: 78.39%; 02: -8.467%; 03: 12.89%; 04: -0.8391%; 05: -0.8391%; 06: 19.76%; 07: -4.195%; 08: 446.2%
| Z-Score (Absolute) |  | 01: 4.07; 02: 0.82; 03: 0.39; 04: 0.39; 05: 0.39; 06: 0.77; 07: 0.58; 08: 24.79
| Analytical results for Te in OREAS 605b (Certified Value 32.2 ppm) | [23 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 31.1289 | 30.5 | 36.28 | 32.5 | 33.3 | 28.5 | 32.2 | 31.9 | 31.4 | 34 | 26.9661 | 28.4 | 32.4 | 28.2 | 38.7 | 32.3 | 35.0299 | 35.0886 |  |  |  |  | 1
| 1 | 2 | 31.0709 | 30.3 | 33.55 | 33.2 | 33.5 | 28.2 | 33.4 | 31.7 | 31.9 | 35.3 | 25.8589 | 26.43 | 32.6 | 23.4 | 36.8 | 32 | 34.798 | 39.5065 |  |  |  |  | 30
| 1 | 3 | 30.9493 | 30.5 | 35.2 | 33.9 | 33 | 28.7 | 31.6 | 31.9 | 30.7 | 35.9 | 28.3402 | 28.92 | 32.3 | 27.4 | 36.9 | 32.9 | 34.5143 | 35.9146 |  |  |  |  | 16
| 1 | 4 | 30.8542 | 29.5 | 35.86 | 33.7 | 33.8 | 28.9 | 36.3 | 30 | 30.7 | 36.8 | 27.5343 | 25.66 | 32.3 | 27.4 | 36.8 | 31.9 | 35.21 | 37.3386 |  |  |  |  | 32.2211
| 1 | 5 | 30.8622 | 29.2 | 34.43 | 33.5 | 31.7 | 30.2 | 36.1 | 31.6 | 31.1 | 34.8 | 24.8557 | 25.65 | 32 | 29.2 | 37.5 | 33.6 | 37.2426 | 37.2112 |  |  |  |  | 59
| 1 | 6 | 30.9128 | 30 | 35.51 | 34.3 | 33.5 | 30.2 | 35.7 | 29.8 | 31.6 | 35.5 | 24.3263 | 26.5 | 33.3 | 28.4 | 38.3 | 32 | 34.5661 | 38.286
| Mean |  | 30.963 | 30 | 35.1383 | 33.5167 | 33.1333 | 29.1167 | 34.2167 | 31.15 | 31.2333 | 35.3833 | 26.3136 | 26.9267 | 32.4833 | 27.3333 | 37.5 | 32.45 | 35.2268 | 37.2242
| Median |  | 30.931 | 30.15 | 35.355 | 33.6 | 33.4 | 28.8 | 34.55 | 31.65 | 31.25 | 35.4 | 26.4125 | 26.465 | 32.35 | 27.8 | 37.2 | 32.15 | 34.9139 | 37.2749
| Std Dev. |  | 0.11303 | 0.54406 | 0.99907 | 0.62102 | 0.75011 | 0.87044 | 2.0817 | 0.97724 | 0.48854 | 0.95376 | 1.5687 | 1.4005 | 0.4446 | 2.0422 | 0.82704 | 0.67157 | 1.0228 | 1.5891
| Rel.Std.Dev. |  | 0.365% | 1.814% | 2.843% | 1.853% | 2.264% | 2.989% | 6.084% | 3.137% | 1.564% | 2.696% | 5.962% | 5.201% | 1.369% | 7.472% | 2.205% | 2.07% | 2.903% | 4.269%
| PDM3 |  | -3.904% | -6.893% | 9.054% | 4.021% | 2.831% | -9.635% | 6.193% | -3.324% | -3.066% | 9.814% | -18.33% | -16.43% | 0.8139% | -15.17% | 16.38% | 0.7105% | 9.328% | 15.53%
| Z-Score (Absolute) |  | 0.39 | 0.65 | 0.7 | 0.28 | 0.17 | 0.88 | 0.46 | 0.35 | 0.32 | 0.77 | 1.61 | 1.45 | 0 | 1.35 | 1.32 | 0 | 0.72 | 1.25
| Analytical results for Th in OREAS 605b (Certified Value 4.79 ppm) | [23 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 6.6305 | 5.9 | 2.5 | 3.44 | 5.14 | 4.99 | 4.63 | 5.31 | 4.34 | 7.5 | 4.0823 | 3 | 3.5 | 6.5 | 6.3 | 3.85 | 10.4517 | < 10 |  |  |  |  | 1
| 1 | 2 | 6.6112 | 6.13 | 1.9 | 3.61 | 5.49 | 5.05 | 4.71 | 5.08 | 4.3 | 7.53 | 4.0396 | 2.9 | 3.9 | 5.9 | 6.3 | 1.97 | 10.3247 | < 10 |  |  |  |  | 15
| 1 | 3 | 6.6052 | 6.33 | 1.9 | 3.22 | 5.2 | 5.22 | 3.37 | 5.05 | 4.2 | 7.23 | 3.7828 | 3.3 | 3.3 | 5.4 | 6.6 | 3.27 | 10.9641 | < 10 |  |  |  |  | 16
| 1 | 4 | 6.4896 | 6.38 | 2 | 3.46 | 5.21 | 5.06 | 5.24 | 5.08 | 4.33 | 7.62 | 3.8364 | 3.5 | 3.2 | 5.8 | 5.9 | 3.47 | 10.5473 | < 10 |  |  |  |  | 4.7941
| 1 | 5 | 6.4782 | 5.86 | 2.1 | 3.55 | 5.25 | 4.85 | 5.38 | 5.2 | 4.28 | 7.34 | 3.5907 | 3.3 | 3.9 | 5 | 6.5 | 3.38 | 11.003 | < 10 |  |  |  |  | 60
| 1 | 6 | 6.6654 | 6.1 | 1.8 | 4.06 | 5.8 | 4.91 | 5.2 | 4.99 | 4.3 | 7.34 | 3.3231 | 3.5 | 4 | 5.9 | 6.3 | 3.66 | 10.6071 | < 10
| Mean |  | 6.58 | 6.1167 | 2.0333 | 3.5567 | 5.3483 | 5.0133 | 4.755 | 5.1183 | 4.2917 | 7.4267 | 3.7758 | 3.25 | 3.6333 | 5.75 | 6.3167 | 3.2667 | 10.6496
| Median |  | 6.6082 | 6.115 | 1.95 | 3.505 | 5.23 | 5.02 | 4.955 | 5.08 | 4.3 | 7.42 | 3.8096 | 3.3 | 3.7 | 5.85 | 6.3 | 3.425 | 10.5772
| Std Dev. |  | 0.077444 | 0.2136 | 0.25033 | 0.28019 | 0.2523 | 0.12972 | 0.74293 | 0.11618 | 0.049967 | 0.14638 | 0.28512 | 0.251 | 0.34448 | 0.50892 | 0.24014 | 0.66791 | 0.27597
| Rel.Std.Dev. |  | 1.177% | 3.492% | 12.31% | 7.878% | 4.717% | 2.587% | 15.62% | 2.27% | 1.164% | 1.971% | 7.551% | 7.723% | 9.481% | 8.851% | 3.802% | 20.45% | 2.591%
| PDM3 |  | 37.25% | 27.59% | -57.59% | -25.81% | 11.56% | 4.572% | -0.8162% | 6.762% | -10.48% | 54.91% | -21.24% | -32.21% | -24.21% | 19.94% | 31.76% | -31.86% | 122.1%
| Z-Score (Absolute) |  | 0.84 | 0.59 | 1.58 | 0.77 | 0.18 | 0 | 0.13 | 0.06 | 0.38 | 1.28 | 0.65 | 0.93 | 0.73 | 0.39 | 0.7 | 0.92 | 3 | 0
| Analytical results for Ti in OREAS 605b (Certified Value 0.212 wt.%) | [23 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.1956 | 0.2225 | 0.223 | 0.2 | 0.213 | 0.213 | 0.213 | 0.219 | 0.215 | 0.234 | 0.2154 | 0.22512 | 0.201 | 0.21 | 0.22 | 0.22 | 0.199 | 0.2 | 0.21119 | 0.20997 | 0.1984 | 0.205 | 1
| 1 | 2 | 0.19026 | 0.2191 | 0.217 | 0.2 | 0.213 | 0.22 | 0.213 | 0.219 | 0.212 | 0.224 | 0.2112 | 0.22104 | 0.224 | 0.2 | 0.22 | 0.21 | 0.198 | 0.2 | 0.2125 | 0.20482 | 0.1964 | 0.199 | 32
| 1 | 3 | 0.18608 | 0.2177 | 0.227 | 0.2 | 0.212 | 0.212 | 0.21 | 0.205 | 0.213 | 0.226 | 0.2103 | 0.22496 | 0.226 | 0.2 | 0.22 | 0.21 | 0.2 | 0.2 | 0.21519 | 0.20802 | 0.1959 | 0.205 | 16
| 1 | 4 | 0.19613 | 0.2184 | 0.226 | 0.2 | 0.217 | 0.21 | 0.211 | 0.228 | 0.214 | 0.223 | 0.2183 | 0.22205 | 0.221 | 0.21 | 0.22 | 0.22 | 0.196 | 0.2 | 0.20325 | 0.20965 | 0.1958 | 0.204 | 0.21154
| 1 | 5 | 0.19567 | 0.21 | 0.227 | 0.2 | 0.214 | 0.215 | 0.212 | 0.231 | 0.209 | 0.228 | 0.2164 | 0.22799 | 0.225 | 0.2 | 0.21 | 0.22 | 0.201 | 0.19 | 0.21177 | 0.20598 | 0.1934 | 0.226 | 61
| 1 | 6 | 0.19606 | 0.2205 | 0.226 | 0.22 | 0.218 | 0.224 | 0.217 | 0.228 | 0.21 | 0.223 | 0.215 | 0.22395 | 0.205 | 0.19 | 0.22 | 0.21 | 0.2 | 0.2 | 0.21025 | 0.20477 | 0.1945 | 0.208
| Mean |  | 0.1933 | 0.21803 | 0.22433 | 0.20333 | 0.2145 | 0.21567 | 0.21267 | 0.22167 | 0.21217 | 0.22633 | 0.21443 | 0.22418 | 0.217 | 0.20167 | 0.21833 | 0.215 | 0.199 | 0.19833 | 0.21069 | 0.2072 | 0.19573 | 0.20783
| Median |  | 0.19564 | 0.21875 | 0.226 | 0.2 | 0.2135 | 0.214 | 0.2125 | 0.2235 | 0.2125 | 0.225 | 0.2152 | 0.22445 | 0.2225 | 0.2 | 0.22 | 0.215 | 0.1995 | 0.2 | 0.21148 | 0.207 | 0.19585 | 0.205
| Std Dev. |  | 0.0041929 | 0.0042884 | 0.0038816 | 0.008165 | 0.002429 | 0.0053166 | 0.0024221 | 0.0095847 | 0.0023166 | 0.0042269 | 0.0030859 | 0.0024666 | 0.011045 | 0.0075277 | 0.0040825 | 0.0054772 | 0.0017889 | 0.0040825 | 0.0040134 | 0.0023414 | 0.0017061 | 0.0093684
| Rel.Std.Dev. |  | 2.169% | 1.967% | 1.73% | 4.016% | 1.132% | 2.465% | 1.139% | 4.324% | 1.092% | 1.868% | 1.439% | 1.1% | 5.09% | 3.733% | 1.87% | 2.548% | 0.8989% | 2.058% | 1.905% | 1.13% | 0.8716% | 4.508%
| PDM3 |  | -8.621% | 3.072% | 6.05% | -3.878% | 1.401% | 1.953% | 0.5346% | 4.789% | 0.2983% | 6.995% | 1.37% | 5.979% | 2.583% | -4.665% | 3.213% | 1.638% | -5.926% | -6.241% | -0.3992% | -2.048% | -7.47% | -1.75%
| Z-Score (Absolute) |  | 2.26 | 0.5 | 1.21 | 1.14 | 0.11 | 0.24 | 0.1 | 0.91 | 0.15 | 1.43 | 0.1 | 1.19 | 0.39 | 1.33 | 0.53 | 0.16 | 1.63 | 1.7 | 0.32 | 0.71 | 1.99 | 0.64
| Analytical results for Tl in OREAS 605b (Certified Value 15.5 ppm) | [23 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 15.1343 | 15.56 | 18.2 | 14.6 | 16.3 | 15 | 17.15 | 16.25 | 14.2 | 15.56 | 16.518 | 15.08 | 14.7 | 16.8 | 17.74 | 15.8 | 15.49 | 13.7646 |  |  |  |  | 1
| 1 | 2 | 15.4359 | 15.26 | 17.91 | 15 | 16 | 15.05 | 17.55 | 15.8 | 14.4 | 15.91 | 15.9387 | 14.88 | 16 | 14.8 | 17.18 | 16.2 | 15.531 | 15.2986 |  |  |  |  | 33
| 1 | 3 | 15.5001 | 15.61 | 18.27 | 14.7 | 15.3 | 15.85 | 16.4 | 15.5 | 14.5 | 15.8 | 15.6862 | 14.98 | 13.8 | 16.2 | 17.38 | 16.5 | 15.807 | 14.4741 |  |  |  |  | 16
| 1 | 4 | 15.322 | 15.6 | 18.04 | 16.45 | 15.9 | 16.15 | 17.9 | 15.05 | 14.3 | 16.32 | 16.2223 | 15.49 | 13.9 | 16.1 | 17.02 | 15.9 | 15.463 | 14.2884 |  |  |  |  | 15.5388
| 1 | 5 | 15.1884 | 14.72 | 17.5 | 15.75 | 16.65 | 15.55 | 17.95 | 15.5 | 14.5 | 15.61 | 16.5577 | 16.05 | 15.5 | 16.9 | 17.47 | 16.6 | 16.137 | 15.7724 |  |  |  |  | 62
| 1 | 6 | 15.3203 | 14.96 | 16.76 | 15.85 | 15.45 | 15.6 | 17.6 | 14.85 | 14.3 | 16.13 | 16.024 | 15.45 | 16 | 17.7 | 17.65 | 16.1 | 15.574 | 17.8354
| Mean |  | 15.3168 | 15.285 | 17.78 | 15.3917 | 15.9333 | 15.5333 | 17.425 | 15.4917 | 14.3667 | 15.8883 | 16.1578 | 15.3217 | 14.9833 | 16.4167 | 17.4067 | 16.1833 | 15.667 | 15.2389
| Median |  | 15.3211 | 15.41 | 17.975 | 15.375 | 15.95 | 15.575 | 17.575 | 15.5 | 14.35 | 15.855 | 16.1231 | 15.265 | 15.1 | 16.5 | 17.425 | 16.15 | 15.5525 | 14.8863
| Std Dev. |  | 0.13972 | 0.37575 | 0.56917 | 0.73717 | 0.50761 | 0.44796 | 0.57858 | 0.50539 | 0.12111 | 0.29607 | 0.34121 | 0.43485 | 0.99883 | 0.9786 | 0.2742 | 0.31885 | 0.26099 | 1.4623
| Rel.Std.Dev. |  | 0.9122% | 2.458% | 3.201% | 4.789% | 3.186% | 2.884% | 3.32% | 3.262% | 0.843% | 1.863% | 2.112% | 2.838% | 6.666% | 5.961% | 1.575% | 1.97% | 1.666% | 9.596%
| PDM3 |  | -1.428% | -1.633% | 14.42% | -0.9467% | 2.539% | -0.03499% | 12.14% | -0.3031% | -7.543% | 2.25% | 3.984% | -1.397% | -3.575% | 5.65% | 12.02% | 4.148% | 0.8252% | -1.93%
| Z-Score (Absolute) |  | 0.55 | 0.61 | 4.23 | 0.4 | 0.65 | 0.13 | 3.54 | 0.21 | 2.4 | 0.56 | 1.08 | 0.54 | 1.2 | 1.59 | 3.51 | 1.13 | 0.13 | 0.7
| Analytical results for Tm in OREAS 605b (Indicative Value 77.2 ppb) | [23 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab |  |  |  |  | 1
| No | No | 01: 03; 02: 11; 03: 14; 05: 16; 06: 23 |  |  |  |  | ppb
|  |  | 01: 4A*OES/MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 5A*MS |  |  |  |  | 1
1
| 1 | 1 | 01: < 100; 02: 90; 03: < 100; 05: 80; 06: < 100 |  |  |  |  | 1
| 1 | 2 | 01: < 100; 02: 90; 03: < 100; 05: 80; 06: < 100 |  |  |  |  | 7
| 1 | 3 | 01: 100; 02: 90; 03: < 100; 05: 80; 06: < 100 |  |  |  |  | 16
| 1 | 4 | 01: < 100; 02: 90; 03: < 100; 05: 80; 06: < 100 |  |  |  |  | 77.2222
| 1 | 5 | 01: < 100; 02: 100; 03: < 100; 05: 80; 06: < 100 |  |  |  |  | 13
| 1 | 6 | 01: < 100; 02: 100; 03: < 100; 05: 80; 06: < 100
| Mean |  | 01: 100; 02: 93.3333; 05: 80
| Median |  | 01: 100; 02: 90; 05: 80
| Std Dev. |  | 02: 5.164; 05: 0
| Rel.Std.Dev. |  | 02: 5.533%; 05: 0%
| PDM3 |  | 01: 29.5%; 02: 20.86%; 05: 3.597%
| Z-Score (Absolute) |  | 01: 0; 02: 2.83; 03: 0.67; 05: 1.75; 06: 0.67
| Analytical results for U in OREAS 605b (Certified Value 3.11 ppm) | [23 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 23 | 24 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 5A*MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 3.3579 | 3.08 | 3.5 | 2.9 | 3.1 | 3.1 | 3.1 | 3.1 | 3.15 | 3.05 | 3.1185 | 2.9 | 2.41 | 3.3 | 3.5 | 3.1 | 2.906 | < 10 |  |  |  |  | 1
| 1 | 2 | 3.3292 | 3.05 | 3.3 | 3 | 3 | 3.1 | 3.2 | 3 | 3.17 | 3.07 | 3.0281 | 3 | 2.77 | 2.9 | 3.4 | 3.19 | 2.905 | < 10 |  |  |  |  | 35
| 1 | 3 | 3.231 | 3.16 | 3.5 | 2.9 | 3.1 | 3.2 | 3 | 2.9 | 3.18 | 3.07 | 3.068 | 3 | 2.44 | 3.2 | 3.4 | 3.26 | 2.967 | < 10 |  |  |  |  | 16
| 1 | 4 | 3.205 | 3.17 | 3.4 | 3.1 | 3 | 3.2 | 3.4 | 2.9 | 3.11 | 3.17 | 3.0644 | 2.7 | 2.37 | 3.3 | 3.4 | 3.21 | 2.883 | < 10 |  |  |  |  | 3.1125
| 1 | 5 | 3.4774 | 3.06 | 3.3 | 3.1 | 3.1 | 3.1 | 3.3 | 2.9 | 3.14 | 3.05 | 3.1185 | 2.6 | 2.68 | 3.4 | 3.5 | 3.29 | 2.995 | < 10 |  |  |  |  | 63
| 1 | 6 | 3.4185 | 3 | 3.4 | 3.1 | 3.2 | 3.2 | 3.4 | 2.9 | 3.13 | 3.14 | 3.0873 | 2.7 | 2.79 | 3.5 | 3.5 | 3.16 | 2.932 | < 10
| Mean |  | 3.3365 | 3.0867 | 3.4 | 3.0167 | 3.0833 | 3.15 | 3.2333 | 2.95 | 3.1467 | 3.0917 | 3.0808 | 2.8167 | 2.5767 | 3.2667 | 3.45 | 3.2017 | 2.9313
| Median |  | 3.3436 | 3.07 | 3.4 | 3.05 | 3.1 | 3.15 | 3.25 | 2.9 | 3.145 | 3.07 | 3.0776 | 2.8 | 2.56 | 3.3 | 3.45 | 3.2 | 2.919
| Std Dev. |  | 0.10539 | 0.066232 | 0.089443 | 0.098319 | 0.075277 | 0.054772 | 0.1633 | 0.083666 | 0.02582 | 0.050761 | 0.034919 | 0.17224 | 0.19117 | 0.20656 | 0.054772 | 0.068532 | 0.04242
| Rel.Std.Dev. |  | 3.159% | 2.146% | 2.631% | 3.259% | 2.441% | 1.739% | 5.05% | 2.836% | 0.8205% | 1.642% | 1.133% | 6.115% | 7.419% | 6.323% | 1.588% | 2.141% | 1.447%
| PDM3 |  | 7.198% | -0.8289% | 9.238% | -3.078% | -0.936% | 1.206% | 3.883% | -5.22% | 1.099% | -0.6683% | -1.018% | -9.504% | -17.21% | 4.954% | 10.84% | 2.866% | -5.82%
| Z-Score (Absolute) |  | 1.12 | 0.17 | 1.45 | 0.53 | 0.18 | 0.16 | 0.59 | 0.87 | 0.14 | 0.14 | 0.2 | 1.56 | 2.8 | 0.76 | 1.71 | 0.43 | 0.97 | 9.69
| Analytical results for V in OREAS 605b (Certified Value 62 ppm) | [23 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 50.503 | 60 | 56 | 63 | 61 | 59 | 62 | 63 | 60 | 68 | 64 | 61.3534 | 57 | 59.6 | 60 | 62 | 59.3 | 65 | 65.5169 | 62.2436 | 65.85 | 64.275 | 1
| 1 | 2 | 50.08 | 60 | 57 | 61 | 61 | 58 | 61 | 63 | 60 | 67 | 63 | 58.9712 | 62 | 60.4 | 60 | 65 | 59.4 | 62 | 64.8722 | 62.1183 | 65.72 | 62.455 | 36
| 1 | 3 | 49.793 | 61 | 56 | 63 | 61 | 59 | 60 | 57 | 61 | 67 | 62 | 58.1509 | 63 | 59.5 | 60 | 61 | 60 | 65 | 65.8361 | 62.5696 | 64.33 | 66.152 | 16
| 1 | 4 | 51.007 | 59 | 58 | 64 | 61 | 59 | 60 | 62 | 60 | 67 | 64 | 57.4139 | 62 | 60.8 | 60 | 64 | 58.2 | 63 | 61.9783 | 61.96 | 65.85 | 63.381 | 61.7898
| 1 | 5 | 49.607 | 56 | 58 | 64 | 63 | 60 | 60 | 64 | 59 | 68 | 63 | 60.5675 | 63 | 58.4 | 65 | 65 | 60.7 | 64 | 64.8281 | 61.9543 | 65.75 | 65.815 | 64
| 1 | 6 | 50.647 | 60 | 57 | 64 | 64 | 62 | 62 | 63 | 59 | 67 | 63 | 57.6708 | 58 | 61.1 | 60 | 61 | 58 | 62 | 65.0486 | 61.8808 | 64.91 | 65.698
| Mean |  | 50.2728 | 59.3333 | 57 | 63.1667 | 61.8333 | 59.5 | 60.8333 | 62 | 59.8333 | 67.3333 | 63.1667 | 59.0213 | 60.8333 | 59.9667 | 60.8333 | 63 | 59.2667 | 63.5 | 64.68 | 62.1211 | 65.4017 | 64.6293
| Median |  | 50.2915 | 60 | 57 | 63.5 | 61 | 59 | 60.5 | 63 | 60 | 67 | 63 | 58.5611 | 62 | 60 | 60 | 63 | 59.35 | 63.5 | 64.9604 | 62.0391 | 65.735 | 64.9865
| Std Dev. |  | 0.53717 | 1.7512 | 0.89443 | 1.169 | 1.3292 | 1.3784 | 0.98319 | 2.5298 | 0.75277 | 0.5164 | 0.75277 | 1.6123 | 2.6394 | 0.99733 | 2.0412 | 1.8974 | 1.0347 | 1.3784 | 1.3806 | 0.25598 | 0.6348 | 1.5019
| Rel.Std.Dev. |  | 1.069% | 2.951% | 1.569% | 1.851% | 2.15% | 2.317% | 1.616% | 4.08% | 1.258% | 0.7669% | 1.192% | 2.732% | 4.339% | 1.663% | 3.355% | 3.012% | 1.746% | 2.171% | 2.135% | 0.4121% | 0.9706% | 2.324%
| PDM3 |  | -18.64% | -3.976% | -7.752% | 2.228% | 0.07038% | -3.706% | -1.548% | 0.3401% | -3.166% | 8.972% | 2.228% | -4.481% | -1.548% | -2.951% | -1.548% | 1.959% | -4.083% | 2.768% | 4.677% | 0.5361% | 5.845% | 4.595%
| Z-Score (Absolute) |  | 4.07 | 0.74 | 1.59 | 0.67 | 0.18 | 0.67 | 0.18 | 0.25 | 0.55 | 2.21 | 0.67 | 0.85 | 0.18 | 0.5 | 0.18 | 0.61 | 0.76 | 0.8 | 1.23 | 0.29 | 1.5 | 1.21
| Analytical results for W in OREAS 605b (Certified Value 30.2 ppm) | [23 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 33.3854 | 31.5 | 30.7 | 29 | 33.1 | 27.5 | 28.3 | 31 | 31.5 | 27.7 | 30.2 | 26.6143 | 29.2 | 30.6 | 30 | 32.8 | 31.2 | 50 | 30.1477 | < 20 | < 5 | 23.305 | 1
| 1 | 2 | 33.5218 | 31.2 | 28.5 | 30 | 32.6 | 28.3 | 29.8 | 32.7 | 30.9 | 27.9 | 30.5 | 25.7628 | 28.7 | 35.1 | 30 | 32.2 | 30.2 | 43 | 30.0212 | < 20 | < 5 | 25.757 | 37
| 1 | 3 | 33.295 | 32.8 | 31.2 | 28 | 31.1 | 28.5 | 28.8 | 30 | 30.5 | 28 | 31.2 | 25.6432 | 30.8 | 30.7 | 29.5 | 32.2 | 31.7 | 41 | 30.6746 | < 20 | < 5 | 25.232 | 16
| 1 | 4 | 33.6796 | 31.9 | 32 | 29 | 30.7 | 28.8 | 28.4 | 33.9 | 30.3 | 28.5 | 31.8 | 26.5804 | 27.7 | 30.3 | 28 | 31.9 | 31.3 | 43 | 30.0421 | < 20 | < 5 | 24.434 | 30.1677
| 1 | 5 | 33.8854 | 30.6 | 32.6 | 29 | 30.2 | 28.8 | 28.4 | 33.5 | 30.3 | 28.1 | 31.1 | 26.203 | 31.1 | 34.1 | 30 | 32.6 | 32.2 | 46 | 31.394 | < 20 | < 5 | 30.877 | 65
| 1 | 6 | 33.8717 | 31.7 | 30.9 | 28 | 31 | 29.1 | 30.4 | 33.8 | 29.3 | 28.3 | 31 | 25.4778 | 30.9 | 34.9 | 30 | 33 | 31.1 | 40 | 30.5965 | < 20 | < 5 | 24.619
| Mean |  | 33.6065 | 31.6167 | 30.9833 | 28.8333 | 31.45 | 28.5 | 29.0167 | 32.4833 | 30.4667 | 28.0833 | 30.9667 | 26.0469 | 29.7333 | 32.6167 | 29.5833 | 32.45 | 31.2833 | 43.8333 | 30.4794 |  |  | 25.704
| Median |  | 33.6007 | 31.6 | 31.05 | 29 | 31.05 | 28.65 | 28.6 | 33.1 | 30.4 | 28.05 | 31.05 | 25.9829 | 30 | 32.4 | 30 | 32.4 | 31.25 | 43 | 30.3721 |  |  | 24.9255
| Std Dev. |  | 0.24762 | 0.73598 | 1.4106 | 0.75277 | 1.1397 | 0.56214 | 0.87731 | 1.6241 | 0.73121 | 0.28577 | 0.56095 | 0.48958 | 1.4038 | 2.3103 | 0.80104 | 0.41833 | 0.66758 | 3.656 | 0.52906 |  |  | 2.6657
| Rel.Std.Dev. |  | 0.7368% | 2.328% | 4.553% | 2.611% | 3.624% | 1.972% | 3.023% | 5% | 2.4% | 1.018% | 1.811% | 1.88% | 4.721% | 7.083% | 2.708% | 1.289% | 2.134% | 8.341% | 1.736% |  |  | 10.37%
| PDM3 |  | 11.4% | 4.803% | 2.704% | -4.423% | 4.251% | -5.528% | -3.815% | 7.676% | 0.9912% | -6.909% | 2.649% | -13.66% | -1.44% | 8.118% | -1.937% | 7.566% | 3.698% | 45.3% | 1.033% |  |  | -14.8%
| Z-Score (Absolute) |  | 1.17 | 0.43 | 0.19 | 0.61 | 0.36 | 0.74 | 0.54 | 0.75 | 0 | 0.89 | 0.18 | 1.65 | 0.28 | 0.8 | 0.33 | 0.74 | 0.3 | 4.99 | 0 | 7.64 | 10.44 | 1.78
| Analytical results for Y in OREAS 605b (Certified Value 5.98 ppm) | [23 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 26 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 7.549 | 6.1 | 5.8 | < 5 | 5.9 | 5.9 | 6.3 | 6.1 | 6.2 | 5.9 | 6.05 | 5.9772 | 5.7 | 6.1 | 5.8 | 6.4 | 5.64 | 5.5 | 5.3204 | 5.8796 | 5.448 |  | 1
| 1 | 2 | 7.5932 | 6 | 6.1 | < 5 | 6.2 | 5.7 | 6.1 | 6.4 | 6.1 | 5.8 | 6.06 | 6.0336 | 6.5 | 6.2 | 5.3 | 6.2 | 5.68 | 5.4 | 5.292 | 5.7828 | 5.319 |  | 38
| 1 | 3 | 7.4227 | 6 | 6.5 | < 5 | 6.1 | 6.1 | 6.2 | 6.1 | 6.1 | 5.8 | 6.02 | 5.8511 | 6.1 | 6.2 | 5.9 | 6.1 | 5.57 | 5.4 | 5.2962 | 6.0127 | 5.67 |  | 16
| 1 | 4 | 7.4788 | 6 | 6.2 | < 5 | 6.2 | 5.9 | 6.4 | 6.4 | 5.9 | 5.7 | 6.19 | 5.9047 | 6.2 | 6.1 | 5.8 | 5.8 | 5.58 | 5.5 | 5.1104 | 6.0043 | 5.657 |  | 5.9755
| 1 | 5 | 7.9443 | 5.8 | 6.2 | < 5 | 5.8 | 6 | 6 | 6.2 | 6.2 | 5.8 | 6.2 | 5.9457 | 6.6 | 5.9 | 6 | 6.3 | 5.74 | 5.7 | 5.2983 | 5.7019 | 5.331 |  | 66
| 1 | 6 | 7.6037 | 6.1 | 5.6 | < 5 | 5.9 | 6.4 | 6.1 | 6.2 | 5.9 | 5.7 | 6.19 | 5.7804 | 5.4 | 6.2 | 6.1 | 6.4 | 5.78 | 5.2 | 5.2542 | 5.8249 | 5.031
| Mean |  | 7.5986 | 6 | 6.0667 |  | 6.0167 | 6 | 6.1833 | 6.2333 | 6.0667 | 5.7833 | 6.1183 | 5.9155 | 6.0833 | 6.1167 | 5.8167 | 6.2 | 5.665 | 5.45 | 5.2619 | 5.8677 | 5.4093
| Median |  | 7.5711 | 6 | 6.15 |  | 6 | 5.95 | 6.15 | 6.2 | 6.1 | 5.8 | 6.125 | 5.9252 | 6.15 | 6.15 | 5.85 | 6.25 | 5.66 | 5.45 | 5.2941 | 5.8523 | 5.3895
| Std Dev. |  | 0.18295 | 0.10954 | 0.32042 |  | 0.17224 | 0.23664 | 0.1472 | 0.13663 | 0.13663 | 0.075277 | 0.083287 | 0.090745 | 0.46224 | 0.1169 | 0.27869 | 0.22804 | 0.084794 | 0.16432 | 0.077266 | 0.12356 | 0.24009
| Rel.Std.Dev. |  | 2.408% | 1.826% | 5.282% |  | 2.863% | 3.944% | 2.381% | 2.192% | 2.252% | 1.302% | 1.361% | 1.534% | 7.598% | 1.911% | 4.791% | 3.678% | 1.497% | 3.015% | 1.468% | 2.106% | 4.438%
| PDM3 |  | 27.16% | 0.4104% | 1.526% |  | 0.6893% | 0.4104% | 3.478% | 4.315% | 1.526% | -3.216% | 2.391% | -1.004% | 1.805% | 2.363% | -2.658% | 3.757% | -5.196% | -8.794% | -11.94% | -1.804% | -9.474%
| Z-Score (Absolute) |  | 5.88 | 0 | 0.25 | 12.87 | 0.06 | 0 | 0.67 | 0.86 | 0.25 | 0.8 | 0.44 | 0.31 | 0.31 | 0.43 | 0.67 | 0.74 | 1.23 | 2.02 | 2.71 | 0.49 | 2.17
| Analytical results for Yb in OREAS 605b (Certified Value 0.6 ppm) | [23 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 01; 02: 03; 03: 11; 04: 14; 05: 16; 06: 20; 07: 21; 08: 23 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES/MS; 08: 5A*MS |  | 2
3
| 1 | 1 | 01: 0.78152; 02: 0.7; 03: 0.59; 04: 0.6; 05: 0.7; 06: 0.6; 07: 0.571; 08: 0.532 |  | 1
| 1 | 2 | 01: 0.79352; 02: 0.6; 03: 0.6; 04: 0.6; 05: 0.65; 06: 0.6; 07: 0.571; 08: 0.536 |  | 39
| 1 | 3 | 01: 0.77452; 02: 0.6; 03: 0.61; 04: 0.6; 05: 0.7; 06: 0.6; 07: 0.515; 08: 0.546 |  | 16
| 1 | 4 | 01: 0.78082; 02: 0.7; 03: 0.66; 04: 0.6; 05: 0.65; 06: 0.6; 07: 0.546; 08: 0.538 |  | 0.59707
| 1 | 5 | 01: 0.75218; 02: 0.5; 03: 0.63; 04: 0.6; 05: 0.7; 06: 0.5; 07: 0.559; 08: 0.542 |  | 67
| 1 | 6 | 01: 0.78831; 02: 0.5; 03: 0.64; 04: 0.6; 05: 0.7; 06: 0.6; 07: 0.536; 08: 0.555
| Mean |  | 01: 0.77848; 02: 0.6; 03: 0.62167; 04: 0.6; 05: 0.68333; 06: 0.58333; 07: 0.54967; 08: 0.5415
| Median |  | 01: 0.78117; 02: 0.6; 03: 0.62; 04: 0.6; 05: 0.7; 06: 0.6; 07: 0.5525; 08: 0.54
| Std Dev. |  | 01: 0.014454; 02: 0.089443; 03: 0.026394; 04: 0; 05: 0.02582; 06: 0.040825; 07: 0.021906; 08: 0.0081915
| Rel.Std.Dev. |  | 01: 1.857%; 02: 14.91%; 03: 4.246%; 04: 0%; 05: 3.779%; 06: 6.999%; 07: 3.985%; 08: 1.513%
| PDM3 |  | 01: 30.38%; 02: 0.4905%; 03: 4.119%; 04: 0.4905%; 05: 14.45%; 06: -2.301%; 07: -7.94%; 08: -9.307%
| Z-Score (Absolute) |  | 01: 3.34; 02: 0; 03: 0.41; 04: 0; 05: 1.56; 06: 0.31; 07: 0.94; 08: 1.1
| Analytical results for Zn in OREAS 605b (Certified Value 0.24 wt.%) | [23 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 16 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.23324 | 0.2461 | 0.23504 | 0.246 | 0.235 | 0.236 | 0.24 | 0.246 | 0.239 | 0.258 | 0.2281 | 0.22769 | 0.2409 | 0.232 | 0.239 | 0.2449 | 0.238 | 0.2353 | 0.247 | 0.24733 | 0.2459 | 0.23999 | 1
| 1 | 2 | 0.23638 | 0.2427 | 0.2388 | 0.246 | 0.238 | 0.237 | 0.24 | 0.242 | 0.235 | 0.247 | 0.2283 | 0.23028 | 0.2452 | 0.233 | 0.24 | 0.2363 | 0.244 | 0.2383 | 0.24492 | 0.24637 | 0.2421 | 0.23225 | 40
| 1 | 3 | 0.23725 | 0.2417 | 0.23934 | 0.246 | 0.234 | 0.238 | 0.239 | 0.237 | 0.241 | 0.25 | 0.2298 | 0.22977 | 0.2475 | 0.234 | 0.236 | 0.2387 | 0.248 | 0.2508 | 0.2484 | 0.24773 | 0.2465 | 0.23976 | 16
| 1 | 4 | 0.24248 | 0.2438 | 0.24286 | 0.246 | 0.247 | 0.237 | 0.238 | 0.239 | 0.229 | 0.25 | 0.2302 | 0.22565 | 0.2432 | 0.237 | 0.242 | 0.2317 | 0.233 | 0.2345 | 0.23988 | 0.24591 | 0.251 | 0.23746 | 0.23999
| 1 | 5 | 0.23543 | 0.2344 | 0.24319 | 0.246 | 0.243 | 0.238 | 0.238 | 0.233 | 0.235 | 0.251 | 0.2318 | 0.22631 | 0.2476 | 0.224 | 0.238 | 0.2381 | 0.243 | 0.2318 | 0.24914 | 0.24523 | 0.2441 | 0.23928 | 68
| 1 | 6 | 0.23591 | 0.246 | 0.23591 | 0.247 | 0.25 | 0.248 | 0.243 | 0.246 | 0.232 | 0.253 | 0.2314 | 0.227 | 0.2422 | 0.224 | 0.239 | 0.2417 | 0.238 | 0.25 | 0.24781 | 0.24622 | 0.243 | 0.23886
| Mean |  | 0.23678 | 0.24245 | 0.23919 | 0.24617 | 0.24117 | 0.239 | 0.23967 | 0.2405 | 0.23517 | 0.2515 | 0.22993 | 0.22778 | 0.24443 | 0.23067 | 0.239 | 0.23857 | 0.24067 | 0.24012 | 0.24619 | 0.24647 | 0.24543 | 0.23793
| Median |  | 0.23615 | 0.24325 | 0.23907 | 0.246 | 0.2405 | 0.2375 | 0.2395 | 0.2405 | 0.235 | 0.2505 | 0.23 | 0.22734 | 0.2442 | 0.2325 | 0.239 | 0.2384 | 0.2405 | 0.2368 | 0.24741 | 0.24629 | 0.245 | 0.23907
| Std Dev. |  | 0.0030962 | 0.0043168 | 0.0033948 | 0.00040825 | 0.0065549 | 0.0044721 | 0.0018619 | 0.0051672 | 0.0044008 | 0.0037283 | 0.0015332 | 0.0018726 | 0.0027933 | 0.0054283 | 0.002 | 0.0045267 | 0.0053541 | 0.0082344 | 0.0034134 | 0.00092345 | 0.0031985 | 0.0029258
| Rel.Std.Dev. |  | 1.308% | 1.781% | 1.419% | 0.1658% | 2.718% | 1.871% | 0.7769% | 2.149% | 1.871% | 1.482% | 0.6668% | 0.8221% | 1.143% | 2.353% | 0.8368% | 1.897% | 2.225% | 3.429% | 1.386% | 0.3747% | 1.303% | 1.23%
| PDM3 |  | -1.337% | 1.024% | -0.3346% | 2.572% | 0.489% | -0.4138% | -0.136% | 0.2112% | -2.011% | 4.795% | -4.192% | -5.089% | 1.85% | -3.886% | -0.4138% | -0.5944% | 0.2807% | 0.0515% | 2.583% | 2.697% | 2.267% | -0.8587%
| Z-Score (Absolute) |  | 0.74 | 0.61 | 0.17 | 1.49 | 0.3 | 0.21 | 0.05 | 0.14 | 1.12 | 2.76 | 2.37 | 2.88 | 1.08 | 2.2 | 0.21 | 0.32 | 0.18 | 0.05 | 1.5 | 1.56 | 1.32 | 0.47
| Analytical results for Zr in OREAS 605b (Certified Value 77 ppm) | [23 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 5A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 0
1
| 1 | 1 | 75.8 | 76.2 | 70 | 75.3 | 75.5 | 71.9 | 83.8 | 86 | 78.3 | 79.9 | 75.8759 | 63 | 89.6 | 88.7 | 72.9 | 73 | 75.008 | 84.4308 | 84.35 | 69.74 |  |  | 1
| 1 | 2 | 74.4 | 76 | 69 | 76.3 | 74.6 | 72.3 | 86.8 | 84.1 | 78.4 | 72.3 | 75.8222 | 68 | 88.3 | 92.4 | 74.3 | 71.3 | 73.397 | 82.5462 | 84.01 | 70.452 |  |  | 41
| 1 | 3 | 76.6 | 77.7 | 68 | 76.3 | 78.6 | 73.1 | 78.8 | 84 | 77.9 | 73.9 | 74.2613 | 69 | 88 | 92.5 | 73.9 | 71.9 | 73.55 | 82.1631 | 84.17 | 75.218 |  |  | 16
| 1 | 4 | 71.4 | 79.2 | 68 | 82.5 | 76.8 | 75.2 | 89.1 | 79.7 | 78.3 | 76.1 | 75.3192 | 69 | 87.3 | 94 | 73.2 | 73.2 | 72.38 | 85.2821 | 84.58 | 72.936 |  |  | 76.9506
| 1 | 5 | 71.5 | 77 | 68 | 84.9 | 76.2 | 76.1 | 87.4 | 83.9 | 78.2 | 75.7 | 75.3651 | 69 | 84.4 | 92.5 | 75.1 | 77.2 | 73.197 | 82.9541 | 84.02 | 71.854 |  |  | 69
| 1 | 6 | 75.3 | 72.9 | 70 | 84.2 | 82.7 | 75.6 | 87.5 | 77.5 | 79.7 | 76 | 74.3444 | 65 | 86.8 | 91.5 | 72.8 | 69.4 | 73.383 | 80.3033 | 84.08 | 71.621
| Mean |  | 74.1667 | 76.5 | 68.8333 | 79.9167 | 77.4 | 74.0333 | 85.5667 | 82.5333 | 78.4667 | 75.65 | 75.1647 | 67.1667 | 87.4 | 91.9333 | 73.7 | 72.6667 | 73.4858 | 82.9466 | 84.2017 | 71.9702
| Median |  | 74.85 | 76.6 | 68.5 | 79.4 | 76.5 | 74.15 | 87.1 | 83.95 | 78.3 | 75.85 | 75.3421 | 68.5 | 87.65 | 92.45 | 73.55 | 72.45 | 73.39 | 82.7502 | 84.125 | 71.7375
| Std Dev. |  | 2.2223 | 2.1109 | 0.98319 | 4.412 | 2.9251 | 1.8173 | 3.742 | 3.2204 | 0.62823 | 2.5564 | 0.70586 | 2.5626 | 1.7539 | 1.7761 | 0.90111 | 2.6105 | 0.85392 | 1.7561 | 0.22391 | 1.944
| Rel.Std.Dev. |  | 2.996% | 2.759% | 1.428% | 5.521% | 3.779% | 2.455% | 4.373% | 3.902% | 0.8006% | 3.379% | 0.9391% | 3.815% | 2.007% | 1.932% | 1.223% | 3.592% | 1.162% | 2.117% | 0.2659% | 2.701%
| PDM3 |  | -3.618% | -0.5856% | -10.55% | 3.854% | 0.5839% | -3.791% | 11.2% | 7.255% | 1.97% | -1.69% | -2.321% | -12.71% | 13.58% | 19.47% | -4.224% | -5.567% | -4.503% | 7.792% | 9.423% | -6.472%
| Z-Score (Absolute) |  | 0.35 | 0.08 | 1.35 | 0.71 | 0.25 | 0.38 | 1.77 | 1.2 | 0.44 | 0.08 | 0.17 | 1.66 | 2.11 | 2.95 | 0.44 | 0.63 | 0.48 | 1.28 | 1.51 | 0.76
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1055: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 605b (Certified Value 1005 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*AAS | AR*OES/AAS | AR*MS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*AAS | AR*AAS | AR*OES/MS | AR*OES/MS | AR*AAS | AR*OES | AR*AAS | AR*OES | AR*OES/MS | AR*AAS | AR*OES | 0
|  |  | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.4g | 0.5g | 0.5g | 01g | 10g | 01g | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0
| 1 | 1 | 985.2588 | 130.33 | > 100 | 1000 | 985 | > 100 | 1020 | 995 | 1025 | > 250 | > 250 | 933.563 | > 800 | 996 | 1092 | 1010 | 1048 | 936 | N.A. | 1074.6429 | > 200 | 1013 | > 500 | 1
| 1 | 2 | 989.9965 | 137.95 | > 100 | 1000 | 990 | > 100 | NSS | 1000 | 1015 | > 250 | > 250 | 985.6355 | > 800 | 995 | 1108 | 982 | 1042 | 932 | N.A. | 1091.2579 | > 200 | 1022 | > 500 | 25
| 1 | 3 | 983.6234 | 119.76 | > 100 | 1000 | 1000 | > 100 | NSS | 998 | 1025 | > 250 | > 250 | 919.219 | > 800 | 999 | 1095 | 1000 | 1047 | 943 | N.A. | 1073.1499 | > 200 | 1008 | > 500 | 16
| 1 | 4 | 990.654 | 116.84 | > 100 | 1000 | 1010 | > 100 | 997 | 994 | 1030 | > 250 | > 250 | 996.0435 | > 800 | 981 | 1133 | 1000 | 1054 | 928 | N.A. | 1085.2931 | > 200 | 1015 | > 500 | 1005.338
| 1 | 5 | 983.7532 | 89.92 | > 100 | 1000 | 1000 | > 100 | 999 | 986 | 1005 | > 250 | > 250 | 1008.9286 | > 800 | 997 | 1139 | 1000 | 1039 | 939 | N.A. | 1083.978 | > 200 | 1005 | > 500 | 71
| 1 | 6 | 980.2569 | 125.62 | > 100 | 1000 | 989 | > 100 | 1020 | 1010 | 1040 | > 250 | > 250 | 1007.9681 | > 800 | 990 | 1134 | 998 | 1050 | 932 | N.A. | 1079.0261 | > 200 | 998.9 | > 500
| Mean |  | 985.5905 | 120.07 |  | 1000 | 995.6667 |  | 1009 | 997.1667 | 1023.3333 |  |  | 975.2263 |  | 993 | 1116.8333 | 998.3333 | 1046.6667 | 935 |  | 1081.2247 |  | 1010.3167
| Median |  | 984.506 | 122.69 |  | 1000 | 995 |  | 1009.5 | 996.5 | 1025 |  |  | 990.8395 |  | 995.5 | 1120.5 | 1000 | 1047.5 | 934 |  | 1081.502 |  | 1010.5
| Std Dev. |  | 4.0213 | 16.585 |  | 0 | 9.3095 |  | 12.7279 | 7.9099 | 12.1106 |  |  | 39.0441 |  | 6.603 | 21.0658 | 9.0701 | 5.4283 | 5.4406 |  | 6.9015 |  | 8.1192
| Rel.Std.Dev. |  | 0.408% | 13.81% |  | 0% | 0.935% |  | 1.261% | 0.7932% | 1.183% |  |  | 4.004% |  | 0.665% | 1.886% | 0.9085% | 0.5186% | 0.5819% |  | 0.6383% |  | 0.8036%
| PDM3 |  | -1.964% | -88.06% |  | -0.531% | -0.962% |  | 0.3643% | -0.8128% | 1.79% |  |  | -2.995% |  | -1.227% | 11.09% | -0.6967% | 4.111% | -6.996% |  | 7.548% |  | 0.4952%
| Z-Score (Absolute) |  | 0.67 | 46.47 | NA | 0.09 | 0.14 | NA | 0.56 | 0.06 | 1.32 | NA | NA | 1.22 | NA | 0.28 | 6.27 | 0 | 2.56 | 3.35 | NA | 4.39 | NA | 0.63 | NA
| Analytical results for Al in OREAS 605b (Certified Value 1.15 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 3
| 1 | 1 | 1.314 | 1.08 | 0.99 | 1.17 | 1.16 | 1.19 | 1.15 | 1.08 | 1.26 | 1.1657 | 1.239 | 1.14 | 0.88 | 1.06 | 1.16 | 1.18 | 1.21 | 1.045 | 0.99849 | 1.19 | 1.125 |  |  | 1
| 1 | 2 | 1.2695 | 1.09 | 0.95 | 1.16 | 1.15 | 1.18 | 1.14 | 1.08 | 1.25 | 1.1794 | 1.2903 | 1.13 | 0.86 | 1.08 | 1.14 | 1.19 | 1.17 | 1.0562 | 1.0279 | 1.172 | 1.153 |  |  | #N/A
| 1 | 3 | 1.2954 | 1.14 | 0.95 | 1.18 | 1.18 | 1.19 | 1.09 | 1.07 | 1.27 | 1.1686 | 1.2046 | 1.12 | 0.84 | 1.05 | 1.16 | 1.18 | 1.17 | 1.0365 | 1.0471 | 1.225 | 1.158 |  |  | 16
| 1 | 4 | 1.2591 | 1.15 | 0.96 | 1.17 | 1.14 | 1.17 | 1.08 | 1.1 | 1.28 | 1.255 | 1.2787 | 1.16 | 0.85 | 1.15 | 1.13 | 1.19 | 1.19 | 1.0332 | 1.0305 | 1.189 | 1.145 |  |  | 1.1518
| 1 | 5 | 1.2321 | 1.11 | 0.96 | 1.17 | 1.13 | 1.16 | 1.01 | 1.1 | 1.26 | 1.1475 | 1.2278 | 1.12 | 0.84 | 1.11 | 1.18 | 1.16 | 1.1 | 1.003 | 1.0196 | 1.171 | 1.159 |  |  | 72
| 1 | 6 | 1.2813 | 1.13 | 0.96 | 1.12 | 1.13 | 1.06 | 1.06 | 1.09 | 1.26 | 1.1548 | 1.2727 | 1.12 | 0.86 | 1.11 | 1.13 | 1.2 | 1.19 | 0.98923 | 1.0285 | 1.21 | 1.177
| Mean |  | 1.2752 | 1.1167 | 0.96167 | 1.1617 | 1.1483 | 1.1583 | 1.0883 | 1.0867 | 1.2633 | 1.1785 | 1.2522 | 1.1317 | 0.855 | 1.0933 | 1.15 | 1.1833 | 1.1717 | 1.0272 | 1.0253 | 1.1928 | 1.1528
| Median |  | 1.2754 | 1.12 | 0.96 | 1.17 | 1.145 | 1.175 | 1.085 | 1.085 | 1.26 | 1.1671 | 1.2559 | 1.125 | 0.855 | 1.095 | 1.15 | 1.185 | 1.18 | 1.0348 | 1.0282 | 1.1895 | 1.1555
| Std Dev. |  | 0.028636 | 0.028048 | 0.01472 | 0.02137 | 0.019408 | 0.049565 | 0.051929 | 0.012111 | 0.010328 | 0.039084 | 0.033521 | 0.016021 | 0.015166 | 0.037238 | 0.02 | 0.013663 | 0.038166 | 0.025727 | 0.015951 | 0.021274 | 0.017233
| Rel.Std.Dev. |  | 2.246% | 2.512% | 1.531% | 1.84% | 1.69% | 4.279% | 4.771% | 1.114% | 0.8175% | 3.316% | 2.677% | 1.416% | 1.774% | 3.406% | 1.739% | 1.155% | 3.257% | 2.505% | 1.556% | 1.783% | 1.495%
| PDM3 |  | 10.72% | -3.052% | -16.51% | 0.8553% | -0.3023% | 0.5659% | -5.511% | -5.656% | 9.682% | 2.317% | 8.714% | -1.749% | -25.77% | -5.077% | -0.1576% | 2.736% | 1.724% | -10.82% | -10.98% | 3.561% | 0.08843%
| Z-Score (Absolute) |  | 1.97 | 0.52 | 2.96 | 0.18 | 0.03 | 0.13 | 0.97 | 1 | 1.78 | 0.45 | 1.61 | 0.29 | 4.64 | 0.89 | 0 | 0.52 | 0.34 | 1.93 | 1.96 | 0.67 | 0.04
| Analytical results for As in OREAS 605b (Certified Value 2377 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 0
| 1 | 1 | 2198.5 | 2212 | 2333.6 | 2149 | 2510 | 2470 | 2690 | 2560 | 2510 | 2500 | 2446 | 2438.082 | 2149.2 | 2570 | 2190 | 2215 | 2380 | 2214 | 2743.525 | 2260.9494 | 2350 | 2358.215 |  | 1
| 1 | 2 | 2155.9 | 2184 | 2327.3 | 2156 | 2490 | 2470 | 2700 | 2520 | 2510 | 2550 | 2509 | 2495.1314 | 2200.3 | 2510 | 2180 | 2203 | 2340 | 2214 | 2738.767 | 2118.9399 | 2406 | 2377.381 |  | 26
| 1 | 3 | 2194.7 | 2185 | 2415.8 | 2152 | 2540 | 2520 | 2790 | 2230 | 2500 | 2520 | 2490 | 2392.2979 | 2440.2 | 2480 | 2160 | 2236 | 2410 | 2238 | 2790.506 | 2269.9274 | 2394 | 2327.15 |  | 16
| 1 | 4 | 2120.6 | 2139 | 2405.8 | 2155 | 2530 | 2450 | 2680 | 2260 | 2570 | 2520 | 2516 | 2530.0712 | 2102 | 2560 | 2310 | 2236 | 2370 | 2162 | 2741.345 | 2174.9823 | 2358 | 2438.511 |  | 2377.1354
| 1 | 5 | 2124.3 | 2126 | 2435.7 | 2154 | 2510 | 2470 | 2750 | 2150 | 2530 | 2470 | 2518 | 2428.871 | 2295.9 | 2490 | 2280 | 2172 | 2280 | 2260 | 2695.537 | 2133.1838 | 2410 | 2347.873 |  | 73
| 1 | 6 | 2156.4 | 2235 | 2388.1 | 2161 | 2450 | 2480 | 2430 | 2210 | 2510 | 2440 | 2393 | 2463.0232 | 2009.6 | 2510 | 2190 | 2187 | 2410 | 2201 | 2775.184 | 2222.8745 | 2414 | 2413.648
| Mean |  | 2158.4 | 2180.1667 | 2384.3833 | 2154.5 | 2505 | 2476.6667 | 2673.3333 | 2321.6667 | 2521.6667 | 2500 | 2478.6667 | 2457.9128 | 2199.5333 | 2520 | 2218.3333 | 2208.1667 | 2365 | 2214.8333 | 2747.4773 | 2196.8096 | 2388.6667 | 2377.1297
| Median |  | 2156.15 | 2184.5 | 2396.95 | 2154.5 | 2510 | 2470 | 2695 | 2245 | 2510 | 2510 | 2499.5 | 2450.5526 | 2174.75 | 2510 | 2190 | 2209 | 2375 | 2214 | 2742.435 | 2198.9284 | 2400 | 2367.798
| Std Dev. |  | 33.2494 | 41.6817 | 44.5653 | 4.0373 | 32.0936 | 23.3809 | 126.2801 | 173.3686 | 25.6255 | 39.4968 | 49.7902 | 49.3157 | 151.8739 | 36.8782 | 61.1283 | 25.9955 | 49.295 | 33.3492 | 33.0031 | 64.3861 | 27.7897 | 41.9833
| Rel.Std.Dev. |  | 1.54% | 1.912% | 1.869% | 0.1874% | 1.281% | 0.944% | 4.724% | 7.467% | 1.016% | 1.58% | 2.009% | 2.006% | 6.905% | 1.463% | 2.756% | 1.177% | 2.084% | 1.506% | 1.201% | 2.931% | 1.163% | 1.766%
| PDM3 |  | -9.202% | -8.286% | 0.3049% | -9.366% | 5.379% | 4.187% | 12.46% | -2.333% | 6.08% | 5.169% | 4.271% | 3.398% | -7.471% | 6.01% | -6.68% | -7.108% | -0.5105% | -6.828% | 15.58% | -7.586% | 0.4851% | -0.0002422%
| Z-Score (Absolute) |  | 1.07 | 0.97 | 0.02 | 1.09 | 0.6 | 0.46 | 1.41 | 0.28 | 0.68 | 0.57 | 0.47 | 0.37 | 0.87 | 0.67 | 0.78 | 0.83 | 0.08 | 0.8 | 1.77 | 0.89 | 0.04 | 0.02
| Analytical results for B in OREAS 605b (Indicative Value 12.1 ppm) | [24 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [9 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | [9 empty] | 1
| 1 | 1 | < 20 | 7 | 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 11 | < 10 | < 10 | 12.375 | [9 empty] | 1
| 1 | 2 | < 20 | 5 | 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 12 | < 10 | < 10 | 12.617 | [9 empty] | 10
| 1 | 3 | < 20 | 5 | 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 9 | < 10 | < 10 | 13.03 | [9 empty] | 16
| 1 | 4 | < 20 | 6 | 20 | 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 13 | < 10 | < 10 | 18.107 | [9 empty] | 12.0768
| 1 | 5 | < 20 | 6 | 20 | 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 10 | < 10 | < 10 | 15.905 | [9 empty] | 16
| 1 | 6 | < 20 | 7 | 20 | 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 13 | < 10 | < 10 | 11.27
| Mean |  | 02: 6; 03: 20; 04: 13.3333; 11: 11.3333; 15: 13.884
| Median |  | 02: 6; 03: 20; 04: 10; 11: 11.5; 15: 12.8235
| Std Dev. |  | 02: 0.89443; 03: 0; 04: 5.7735; 11: 1.633; 15: 2.5831
| Rel.Std.Dev. |  | 02: 14.91%; 03: 0%; 04: 43.3%; 11: 14.41%; 15: 18.61%
| PDM3 |  | 02: -50.32%; 03: 65.61%; 04: 10.4%; 11: -6.156%; 15: 14.96%
| Z-Score (Absolute) |  | 6.07 | 0.67 | 19.55 | 4.94 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 6.07 | 7.87 | 0.67 | 0.67 | 11.31
| Analytical results for Ba in OREAS 605b (Indicative Value 16.7 ppm) | [24 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  |  | 1
| 1 | 1 | 31 | 6.5 | 11 | 10 | 10 | 10 | 10 | 10 | 88 | 3 | 66.4049 | 4.9 | 15 | 46 | 19.2 | 27 | 124.075 | 24.911 | 1693 | 29.133 |  |  |  | 1
| 1 | 2 | 43 | 5.3 | 10 | 10 | 10 | 10 | 10 | 10 | 87 | 3 | 63.1139 | 4.3 | 15 | 50 | 15.5 | 25 | 121.879 | 25.0391 | 1611 | 24.024 |  |  |  | 6
| 1 | 3 | 38 | 5.6 | 10 | 10 | 10 | 10 | 10 | 10 | 87 | 3 | 68.2112 | 4.8 | 16 | 47 | 12.3 | 25 | 130.269 | 24.3956 | 1696 | 29.088 |  |  |  | 16
| 1 | 4 | 33 | 5.6 | 10 | 10 | 10 | 10 | 10 | 10 | 86 | 4 | 62.8842 | 4.4 | 17 | 45 | 18.8 | 25 | 126.469 | 24.5265 | 1657 | 31.936 |  |  |  | 16.7432
| 1 | 5 | 29 | 5.4 | 10 | 10 | 10 | 10 | 10 | 10 | 86 | 3 | 67.8645 | 4.3 | 15 | 46 | 20.2 | 27 | 135.962 | 27.7006 | 1667 | 32.926 |  |  |  | 17
| 1 | 6 | 35 | 6.2 | 10 | 10 | 10 | 10 | 10 | 10 | 88 | 3 | 67.1573 | 4.3 | 15 | 48 | 17.3 | 24 | 135.294 | 25.1882 | 1622 | 29.464
| Mean |  | 34.8333 | 5.7667 | 10.1667 | 10 | 10 | 10 | 10 | 10 | 87 | 3.1667 | 65.9393 | 4.5 | 15.5 | 47 | 17.2167 | 25.5 | 128.9913 | 25.2935 | 1657.6667 | 29.4285
| Median |  | 34 | 5.6 | 10 | 10 | 10 | 10 | 10 | 10 | 87 | 3 | 66.7811 | 4.35 | 15 | 46.5 | 18.05 | 25 | 128.369 | 24.9751 | 1662 | 29.2985
| Std Dev. |  | 5.0761 | 0.4761 | 0.40825 | 0 | 0 | 0 | 0 | 0 | 0.89443 | 0.40825 | 2.3616 | 0.27568 | 0.83666 | 1.7889 | 2.9144 | 1.2247 | 5.8492 | 1.2174 | 35.3704 | 3.0962
| Rel.Std.Dev. |  | 14.57% | 8.256% | 4.016% | 0% | 0% | 0% | 0% | 0% | 1.028% | 12.89% | 3.581% | 6.126% | 5.398% | 3.806% | 16.93% | 4.803% | 4.535% | 4.813% | 2.134% | 10.52%
| PDM3 |  | 108% | -65.56% | -39.28% | -40.27% | -40.27% | -40.27% | -40.27% | -40.27% | 419.6% | -81.09% | 293.8% | -73.12% | -7.425% | 180.7% | 2.828% | 52.3% | 670.4% | 51.07% | 9801% | 75.76%
| Z-Score (Absolute) |  | 1.26 | 0.72 | 0.42 | 0.43 | 0.43 | 0.43 | 0.43 | 0.43 | 4.83 | 0.9 | 3.39 | 0.81 | 0.06 | 2.09 | 0.06 | 0.62 | 7.7 | 0.61 | 112.17 | 0.89
| Analytical results for Be in OREAS 605b (Certified Value 0.15 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 25 | 26 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 01g | 0.25g |  |  |  | 2
| 1 | 1 | < 0.5 | 0.3 | < 5 | 0.18 | 0.16 | 0.14 | 0.19 | 0.15 | 0.14 | 0.16 | < 0.5 | 0.1 | 0.16 | < 0.5 | 0.1 | 0.144 | < 0.5 | 0.109 | < 1 | < 0.5 |  |  |  | 1
| 1 | 2 | < 0.5 | < 0.1 | < 5 | 0.15 | 0.16 | 0.14 | 0.17 | 0.16 | 0.14 | 0.17 | < 0.5 | 0.2 | 0.17 | < 0.5 | 0.1 | < 0.1 | < 0.5 | 0.109 | < 1 | < 0.5 |  |  |  | 28
| 1 | 3 | < 0.5 | < 0.1 | < 5 | 0.19 | 0.12 | 0.17 | 0.12 | 0.12 | 0.15 | 0.17 | < 0.5 | 0.2 | 0.14 | < 0.5 | 0.1 | 0.172 | < 0.5 | 0.104 | < 1 | < 0.5 |  |  |  | 16
| 1 | 4 | < 0.5 | < 0.1 | < 5 | 0.23 | 0.29 | 0.14 | 0.13 | 0.14 | 0.15 | 0.18 | < 0.5 | 0.1 | 0.15 | < 0.5 | 0.2 | 0.147 | < 0.5 | 0.099 | < 1 | < 0.5 |  |  |  | 0.15289
| 1 | 5 | < 0.5 | 0.2 | < 5 | 0.17 | 0.19 | 0.15 | 0.16 | 0.15 | 0.14 | 0.15 | < 0.5 | 0.2 | 0.16 | < 0.5 | 0.1 | 0.188 | < 0.5 | 0.106 | < 1 | < 0.5 |  |  |  | 74
| 1 | 6 | < 0.5 | 0.2 | < 5 | 0.17 | 0.27 | 0.16 | 0.14 | 0.16 | 0.15 | 0.17 | < 0.5 | 0.2 | 0.16 | < 0.5 | 0.1 | 0.154 | < 0.5 | 0.107 | < 1 | < 0.5
| Mean |  |  | 0.23333 |  | 0.18167 | 0.19833 | 0.15 | 0.15167 | 0.14667 | 0.145 | 0.16667 |  | 0.16667 | 0.15667 |  | 0.11667 | 0.161 |  | 0.10567
| Median |  |  | 0.2 |  | 0.175 | 0.175 | 0.145 | 0.15 | 0.15 | 0.145 | 0.17 |  | 0.2 | 0.16 |  | 0.1 | 0.154 |  | 0.1065
| Std Dev. |  |  | 0.057735 |  | 0.027142 | 0.067355 | 0.012649 | 0.026394 | 0.015055 | 0.0054772 | 0.010328 |  | 0.05164 | 0.010328 |  | 0.040825 | 0.018601 |  | 0.0037771
| Rel.Std.Dev. |  |  | 24.74% |  | 14.94% | 33.96% | 8.433% | 17.4% | 10.27% | 3.777% | 6.197% |  | 30.98% | 6.592% |  | 34.99% | 11.55% |  | 3.575%
| PDM3 |  |  | 52.62% |  | 18.82% | 29.72% | -1.89% | -0.7994% | -4.07% | -5.16% | 9.012% |  | 9.012% | 2.471% |  | -23.69% | 5.305% |  | -30.89%
| Z-Score (Absolute) |  | 1.77 | 0.53 | 49.69 | 0.32 | 0.67 | 0.35 | 0.32 | 0.43 | 0.46 | 0 | 1.77 | NA | 0.21 | 1.77 | 1.06 | 0.51 | 1.77 | 1.3 | 7.1 | 1.77
Results from laboratory 14 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 605b (Certified Value 50 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.25g | 0.2g | 01g | 0.25g |  |  | 1
| 1 | 1 | 51.87 | 58.51 | 12 | 48.9 | 48.9 | 48.8 | 55.8 | 54.1 | 51.7 | 53.73 | 48.0179 | 42.65 | 49.9 | 54 | 49.32 | 49.8 | 56 | 51.219 | 28.0504 | < 5 | 50.671 |  |  | 1
| 1 | 2 | 48.13 | 56.18 | 12 | 48.2 | 49.7 | 47.9 | 53.9 | 53.5 | 52.6 | 59.18 | 50.2769 | 42.47 | 48.9 | 43 | 49.47 | 49.4 | 58 | 53.266 | 53.9623 | < 5 | 50.101 |  |  | 29
| 1 | 3 | 51.24 | 57.08 | 11 | 49.4 | 50.2 | 50.2 | 48.9 | 52 | 52.3 | 56.93 | 49.5271 | 44.22 | 49.2 | 47 | 49.29 | 49.8 | 57 | 53.636 | 42.2505 | < 5 | 48.587 |  |  | 16
| 1 | 4 | 48.53 | 56.97 | 11 | 49.1 | 46.8 | 50.1 | 45.7 | 54.2 | 52.7 | 53.77 | 51.4883 | 43.65 | 49.4 | 49 | 50.14 | 48.9 | 58 | 53.641 | 57.9038 | < 5 | 51.379 |  |  | 50.2386
| 1 | 5 | 51.74 | 56.01 | 11 | 48 | 46.1 | 45.3 | 49.7 | 52.3 | 52.6 | 57.48 | 49.2845 | 44.45 | 48.3 | 50 | 49.09 | 48.4 | 58 | 53.418 | 33.3304 | < 5 | 50.986 |  |  | 75
| 1 | 6 | 47.89 | 55.59 | 10 | 48.4 | 49 | 46.9 | 51 | 51 | 52.1 | 53.92 | 50.2554 | 44.97 | 50.1 | 58 | 49.24 | 49.6 | 53 | 51.632 | 50.6145 | < 5 | 46.216
| Mean |  | 49.9 | 56.7233 | 11.1667 | 48.6667 | 48.45 | 48.2 | 50.8333 | 52.85 | 52.3333 | 55.835 | 49.8084 | 43.735 | 49.3 | 50.1667 | 49.425 | 49.3167 | 56.6667 | 52.802 | 44.352 |  | 49.6567
| Median |  | 49.885 | 56.575 | 11 | 48.65 | 48.95 | 48.35 | 50.35 | 52.9 | 52.45 | 55.425 | 49.8913 | 43.935 | 49.3 | 49.5 | 49.305 | 49.5 | 57.5 | 53.342 | 46.4325 |  | 50.386
| Std Dev. |  | 1.9033 | 1.0458 | 0.75277 | 0.55015 | 1.6355 | 1.9058 | 3.6187 | 1.2849 | 0.38297 | 2.3434 | 1.1663 | 1.0056 | 0.6603 | 5.2694 | 0.3712 | 0.56006 | 1.9664 | 1.0834 | 11.89 |  | 1.9464
| Rel.Std.Dev. |  | 3.814% | 1.844% | 6.741% | 1.13% | 3.376% | 3.954% | 7.119% | 2.431% | 0.7318% | 4.197% | 2.342% | 2.299% | 1.339% | 10.5% | 0.751% | 1.136% | 3.47% | 2.052% | 26.81% |  | 3.92%
| PDM3 |  | -0.6741% | 12.91% | -77.77% | -3.129% | -3.56% | -4.058% | 1.184% | 5.198% | 4.169% | 11.14% | -0.8565% | -12.95% | -1.868% | -0.1433% | -1.62% | -1.835% | 12.79% | 5.102% | -11.72% |  | -1.158%
| Z-Score (Absolute) |  | 0.11 | 3.27 | 17.82 | 0.46 | 0.56 | 0.67 | 0.54 | 1.48 | 1.24 | 2.86 | 0.07 | 2.74 | 0.17 | 0.24 | 0.11 | 0.16 | 3.25 | 1.46 | 2.46 | 21.83 | 0
| Analytical results for Ca in OREAS 605b (Certified Value 0.197 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 3
| 1 | 1 | 0.19605 | 0.21 | 0.18 | 0.18 | 0.21 | 0.2 | 0.21 | 0.2 | 0.19 | 0.18 | 0.2 | 0.18106 | 0.2 | 0.18 | 0.23 | 0.19 | 0.198 | 0.2 | 0.18754 | 0.21459 | 0.2118 | 0.202 |  | 1
| 1 | 2 | 0.19998 | 0.21 | 0.19 | 0.18 | 0.21 | 0.2 | 0.2 | 0.2 | 0.19 | 0.19 | 0.21 | 0.19678 | 0.2 | 0.18 | 0.21 | 0.19 | 0.19 | 0.2 | 0.18784 | 0.21729 | 0.2099 | 0.204 |  | #N/A
| 1 | 3 | 0.19877 | 0.21 | 0.2 | 0.18 | 0.21 | 0.21 | 0.21 | 0.19 | 0.19 | 0.19 | 0.2 | 0.18607 | 0.2 | 0.18 | 0.2 | 0.2 | 0.191 | 0.19 | 0.18317 | 0.22066 | 0.203 | 0.203 |  | 16
| 1 | 4 | 0.20119 | 0.2 | 0.2 | 0.18 | 0.21 | 0.21 | 0.2 | 0.18 | 0.19 | 0.19 | 0.21 | 0.19425 | 0.2 | 0.18 | 0.24 | 0.19 | 0.196 | 0.19 | 0.18742 | 0.21587 | 0.2175 | 0.202 |  | 0.19695
| 1 | 5 | 0.20289 | 0.2 | 0.2 | 0.17 | 0.21 | 0.2 | 0.2 | 0.17 | 0.19 | 0.19 | 0.2 | 0.18668 | 0.2 | 0.17 | 0.21 | 0.2 | 0.196 | 0.19 | 0.1853 | 0.221 | 0.2069 | 0.205 |  | 76
| 1 | 6 | 0.20874 | 0.2 | 0.19 | 0.18 | 0.2 | 0.2 | 0.19 | 0.18 | 0.19 | 0.18 | 0.19 | 0.19402 | 0.2 | 0.17 | 0.25 | 0.19 | 0.195 | 0.19 | 0.18597 | 0.21929 | 0.2034 | 0.206
| Mean |  | 0.20127 | 0.205 | 0.19333 | 0.17833 | 0.20833 | 0.20333 | 0.20167 | 0.18667 | 0.19 | 0.18667 | 0.20167 | 0.18981 | 0.2 | 0.17667 | 0.22333 | 0.19333 | 0.19433 | 0.19333 | 0.18621 | 0.21812 | 0.20875 | 0.20367
| Median |  | 0.20059 | 0.205 | 0.195 | 0.18 | 0.21 | 0.2 | 0.2 | 0.185 | 0.19 | 0.19 | 0.2 | 0.19035 | 0.2 | 0.18 | 0.22 | 0.19 | 0.1955 | 0.19 | 0.18669 | 0.21829 | 0.2084 | 0.2035
| Std Dev. |  | 0.0043265 | 0.0054772 | 0.008165 | 0.0040825 | 0.0040825 | 0.005164 | 0.0075277 | 0.012111 | 3.0405e-17 | 0.005164 | 0.0075277 | 0.0061046 | 3.0405e-17 | 0.005164 | 0.019664 | 0.005164 | 0.0031411 | 0.005164 | 0.0017894 | 0.0026213 | 0.005519 | 0.001633
| Rel.Std.Dev. |  | 2.15% | 2.672% | 4.223% | 2.289% | 1.96% | 2.54% | 3.733% | 6.488% | 1.6e-14% | 2.766% | 3.733% | 3.216% | 1.52e-14% | 2.923% | 8.805% | 2.671% | 1.616% | 2.671% | 0.961% | 1.202% | 2.644% | 0.8018%
| PDM3 |  | 2.191% | 4.085% | -1.838% | -9.454% | 5.778% | 3.239% | 2.393% | -5.223% | -3.531% | -5.223% | 2.393% | -3.627% | 1.547% | -10.3% | 13.39% | -1.838% | -1.331% | -1.838% | -5.457% | 10.74% | 5.989% | 3.408%
| Z-Score (Absolute) |  | 0.38 | 0.73 | 0.36 | 1.75 | 1.04 | 0.57 | 0.42 | 0.97 | 0.67 | 0.97 | 0.42 | 0.68 | 0.26 | 1.9 | 2.43 | 0.36 | 0.26 | 0.36 | 1.02 | 1.95 | 1.08 | 0.6
| Analytical results for Cd in OREAS 605b (Certified Value 14 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 2
| 1 | 1 | 13.225 | 11.14 | 13.73 | 20 | 13.25 | 13.35 | 14.6 | 14.65 | 13.55 | 11.1 | 13.8 | 13.6321 | 11.12 | 14.6 | 14 | 14.73 | 13.4 | 13 | 16.694 | 13.5907 | 39.19 | 15.584 |  | 1
| 1 | 2 | 13.159 | 11.36 | 13.95 | 21 | 13.5 | 13.6 | 13.9 | 14.1 | 13.4 | 10.9 | 14.14 | 13.7704 | 11.65 | 14.5 | 14 | 14.4 | 13.2 | 12 | 16.48 | 13.4318 | 40.23 | 16.229 |  | 30
| 1 | 3 | 13.096 | 11.43 | 14.28 | 20 | 13.55 | 13.7 | 15.25 | 14.15 | 13.45 | 11 | 13.91 | 13.6817 | 11.45 | 14.6 | 14 | 14.93 | 13.3 | 13 | 16.212 | 13.7066 | 40.21 | 15.349 |  | 16
| 1 | 4 | 13.528 | 11.1 | 14.67 | 21 | 13.25 | 13.35 | 14.1 | 13.2 | 13.65 | 11.2 | 14.59 | 13.6909 | 10.49 | 14.7 | 15 | 14.98 | 13.5 | 13 | 16.302 | 13.4044 | 40.47 | 15.862 |  | 13.9828
| 1 | 5 | 12.976 | 11.33 | 14.43 | 19 | 13.65 | 12.85 | 14.75 | 14.25 | 12.75 | 10.9 | 14.33 | 13.6319 | 11.61 | 14.2 | 15 | 14.39 | 13.3 | 13 | 15.972 | 13.8413 | 39.91 | 15.551 |  | 77
| 1 | 6 | 14.147 | 11.52 | 14.49 | 20 | 13.1 | 13.75 | 13.8 | 14.1 | 13.05 | 11 | 13.31 | 13.7277 | 11.86 | 14.5 | 14 | 14.46 | 13.6 | 12 | 15.907 | 13.6341 | 39.53 | 16.036
| Mean |  | 13.3552 | 11.3133 | 14.2583 | 20.1667 | 13.3833 | 13.4333 | 14.4 | 14.075 | 13.3083 | 11.0167 | 14.0133 | 13.6891 | 11.3633 | 14.5167 | 14.3333 | 14.6483 | 13.3833 | 12.6667 | 16.2612 | 13.6015 | 39.9233 | 15.7685
| Median |  | 13.192 | 11.345 | 14.355 | 20 | 13.375 | 13.475 | 14.35 | 14.125 | 13.425 | 11 | 14.025 | 13.6863 | 11.53 | 14.55 | 14 | 14.595 | 13.35 | 13 | 16.257 | 13.6124 | 40.06 | 15.723
| Std Dev. |  | 0.4297 | 0.16391 | 0.3542 | 0.75277 | 0.2137 | 0.33267 | 0.56303 | 0.47618 | 0.3412 | 0.1169 | 0.44715 | 0.054176 | 0.49379 | 0.17224 | 0.5164 | 0.26828 | 0.1472 | 0.5164 | 0.29946 | 0.16577 | 0.48327 | 0.33157
| Rel.Std.Dev. |  | 3.218% | 1.449% | 2.484% | 3.733% | 1.597% | 2.476% | 3.91% | 3.383% | 2.564% | 1.061% | 3.191% | 0.3958% | 4.345% | 1.186% | 3.603% | 1.832% | 1.1% | 4.077% | 1.842% | 1.219% | 1.21% | 2.103%
| PDM3 |  | -4.488% | -19.09% | 1.971% | 44.23% | -4.287% | -3.929% | 2.984% | 0.6596% | -4.823% | -21.21% | 0.2186% | -2.1% | -18.73% | 3.818% | 2.507% | 4.76% | -4.287% | -9.412% | 16.29% | -2.727% | 185.5% | 12.77%
| Z-Score (Absolute) |  | 0.4 | 2.81 | 0.67 | NA | 0.36 | 0.3 | 0.84 | 0.46 | 0.45 | 3.17 | 0.38 | 0 | 2.76 | 0.98 | NA | 1.14 | 0.36 | NA | 3.05 | 0.1 | 31.08 | 2.46
Results from laboratories 4, 19 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 605b (Certified Value 10.9 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | [6 empty] | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 1
| 1 | 1 | < 10 | 16.18 | 9.7 | 10.8 | 9.22 | 11.6 | 9.84 | 10.15 | 15.5 | 9.703 | 14.5113 | 9.9 | 13.3 | 12.03 | 12.1 | 12.679 | 10.3586 | [6 empty] | 1
| 1 | 2 | < 10 | 16.31 | 9.5 | 10.6 | 9.65 | 9.96 | 9.87 | 10.15 | 15.7 | 10.258 | 14.8527 | 9.6 | 12.9 | 11.82 | 11.7 | 12.633 | 10.9741 | [6 empty] | 31
| 1 | 3 | < 10 | 16.27 | 10.4 | 10.8 | 10.05 | 10.75 | 10.3 | 10.75 | 15.8 | 9.678 | 14.7776 | 10.2 | 13.6 | 11.89 | 11.4 | 12.85 | 10.3025 | [6 empty] | 16
| 1 | 4 | < 10 | 15.63 | 10.1 | 10.5 | 9.27 | 10.7 | 9.67 | 11.7 | 15.8 | 10.24 | 14.5062 | 9.9 | 13.1 | 11.97 | 12.8 | 12.238 | 11.1329 | [6 empty] | 10.9036
| 1 | 5 | < 10 | 15.72 | 9.9 | 10.85 | 9.15 | 10.75 | 10.6 | 9.48 | 15.6 | 9.615 | 14.5473 | 10.1 | 12.8 | 11.91 | 12.7 | 12.473 | 10.8201 | [6 empty] | 78
| 1 | 6 | < 10 | 16.53 | 9.9 | 10.4 | 9.8 | 10.15 | 10.2 | 10.55 | 15.6 | 9.842 | 14.3298 | 10.4 | 13.2 | 12.13 | 12.1 | 12.596 | 10.7777
| Mean |  |  | 16.1067 | 9.9167 | 10.6583 | 9.5233 | 10.6517 | 10.08 | 10.4633 | 15.6667 | 9.8893 | 14.5875 | 10.0167 | 13.15 | 11.9583 | 12.1333 | 12.5782 | 10.7276
| Median |  |  | 16.225 | 9.9 | 10.7 | 9.46 | 10.725 | 10.035 | 10.35 | 15.65 | 9.7725 | 14.5293 | 10 | 13.15 | 11.94 | 12.1 | 12.6145 | 10.7989
| Std Dev. |  |  | 0.35478 | 0.31252 | 0.18552 | 0.36484 | 0.57464 | 0.34728 | 0.74597 | 0.12111 | 0.28836 | 0.19335 | 0.27869 | 0.2881 | 0.11035 | 0.5465 | 0.20696 | 0.33256
| Rel.Std.Dev. |  |  | 2.203% | 3.151% | 1.741% | 3.831% | 5.395% | 3.445% | 7.129% | 0.773% | 2.916% | 1.325% | 2.782% | 2.191% | 0.9228% | 4.504% | 1.645% | 3.1%
| PDM3 |  |  | 47.72% | -9.051% | -2.249% | -12.66% | -2.311% | -7.553% | -4.038% | 43.68% | -9.302% | 33.79% | -8.134% | 20.6% | 9.673% | 11.28% | 15.36% | -1.614%
| Z-Score (Absolute) |  | 3.36 | 3.24 | 0.44 | 0 | 0.67 | 0 | 0.34 | 0.12 | 2.98 | 0.46 | 2.33 | 0.38 | 1.48 | 0.77 | 0.88 | 1.14 | 0.04
| Analytical results for Co in OREAS 605b (Certified Value 18.5 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 2
| 1 | 1 | 16.0348 | 15.9 | 18.3 | 17 | 18.4 | 17.9 | 20.1 | 20.1 | 19.3 | 18.6 | 20.1 | 19.5185 | 18.9 | 18.6 | 15 | 17.9 | 17.6 | 17 | 16.648 | 17.8023 | 17.33 | 18.634 |  | 1
| 1 | 2 | 16.1638 | 15.6 | 18.5 | 17 | 18.8 | 18.4 | 19 | 19.3 | 18.4 | 19.1 | 20.3 | 20.1585 | 18.7 | 18 | 15 | 18 | 17.3 | 17 | 16.572 | 17.8426 | 17.88 | 18.713 |  | 32
| 1 | 3 | 15.9544 | 15.5 | 19.4 | 17 | 18.9 | 17.8 | 20.1 | 19.9 | 18.9 | 19.1 | 20.3 | 19.8534 | 19.9 | 18.1 | 15 | 18.3 | 16.9 | 18 | 16.752 | 17.9569 | 17.25 | 18.257 |  | 16
| 1 | 4 | 16.2866 | 15.2 | 19.8 | 17 | 18.6 | 17.4 | 19.2 | 19 | 18.4 | 19.2 | 20.4 | 20.593 | 18.5 | 18.3 | 18 | 18.2 | 17.2 | 18 | 16.976 | 17.5912 | 17.91 | 19.367 |  | 18.4583
| 1 | 5 | 15.9717 | 15.3 | 20.5 | 16 | 18.8 | 17.2 | 19.1 | 20.9 | 18 | 18.8 | 20.5 | 19.9434 | 19.7 | 17.8 | 16 | 17.8 | 17.2 | 18 | 16.892 | 18.1982 | 18.18 | 18.996 |  | 79
| 1 | 6 | 16.1117 | 16.1 | 20.3 | 17 | 18.3 | 18.3 | 18.5 | 20.2 | 17.4 | 18.7 | 19.1 | 20.3632 | 16.8 | 18.1 | 15 | 18 | 17.6 | 17 | 15.972 | 17.8393 | 17.56 | 18.408
| Mean |  | 16.0871 | 15.6 | 19.4667 | 16.8333 | 18.6333 | 17.8333 | 19.3333 | 19.9 | 18.4 | 18.9167 | 20.1167 | 20.0717 | 18.75 | 18.15 | 15.6667 | 18.0333 | 17.3 | 17.5 | 16.6353 | 17.8717 | 17.685 | 18.7292
| Median |  | 16.0732 | 15.55 | 19.6 | 17 | 18.7 | 17.85 | 19.15 | 20 | 18.4 | 18.95 | 20.3 | 20.0509 | 18.8 | 18.1 | 15 | 18 | 17.25 | 17.5 | 16.7 | 17.8409 | 17.72 | 18.6735
| Std Dev. |  | 0.12643 | 0.34641 | 0.9136 | 0.40825 | 0.24221 | 0.4761 | 0.64083 | 0.67823 | 0.66633 | 0.24833 | 0.51543 | 0.38328 | 1.1059 | 0.27386 | 1.2111 | 0.18619 | 0.26833 | 0.54772 | 0.35769 | 0.19962 | 0.36457 | 0.40335
| Rel.Std.Dev. |  | 0.7859% | 2.221% | 4.693% | 2.425% | 1.3% | 2.67% | 3.315% | 3.408% | 3.621% | 1.313% | 2.562% | 1.91% | 5.898% | 1.509% | 7.73% | 1.032% | 1.551% | 3.13% | 2.15% | 1.117% | 2.061% | 2.154%
| PDM3 |  | -12.85% | -15.49% | 5.463% | -8.803% | 0.9483% | -3.386% | 4.741% | 7.811% | -0.3158% | 2.483% | 8.984% | 8.741% | 1.58% | -1.67% | -15.12% | -2.302% | -6.275% | -5.192% | -9.876% | -3.178% | -4.189% | 1.467%
| Z-Score (Absolute) |  | 2.18 | 2.64 | 1.01 | NA | 0.22 | 0.53 | 0.88 | 1.41 | 0 | 0.49 | 1.62 | 1.58 | 0.33 | 0.24 | NA | 0.35 | 1.04 | NA | 1.66 | 0.5 | 0.67 | 0.31
Results from laboratories 4, 19 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 605b (Certified Value 23.8 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 26 |  |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 0.25g |  |  | 2
| 1 | 1 | 22.236 | 25 | 23.3 | 34 | 24 | 24 | 26 | 25 | 24 | 23 | 27 | 25.4547 | 27.3 | 21.1 | 31 | 22 | 26.6 | 20 | 25.061 | 19.9517 | 22.296 |  |  | 1
| 1 | 2 | 21.7858 | 25 | 22.9 | 33 | 24 | 24 | 25 | 25 | 24 | 22 | 28 | 25.6244 | 28.4 | 21.2 | 32 | 23 | 26.6 | 20 | 25.144 | 19.8879 | 22.516 |  |  | 33
| 1 | 3 | 21.7597 | 25 | 24.7 | 35 | 25 | 24 | 26 | 24 | 23 | 22 | 28 | 25.0486 | 29.1 | 20.6 | 30 | 22 | 27.3 | 21 | 24.672 | 20.5671 | 22.076 |  |  | 16
| 1 | 4 | 22.0633 | 25 | 24.8 | 35 | 24 | 23 | 26 | 24 | 24 | 22 | 29 | 26.4168 | 28.8 | 20.5 | 34 | 21 | 26.7 | 21 | 24.905 | 20.5881 | 23.106 |  |  | 23.8484
| 1 | 5 | 22.4355 | 23 | 25 | 34 | 24 | 23 | 25 | 22 | 24 | 23 | 27 | 25.8156 | 27.3 | 21 | 32 | 22 | 26.6 | 20 | 24.499 | 21.4083 | 22.701 |  |  | 80
| 1 | 6 | 22.357 | 25 | 24.7 | 35 | 24 | 24 | 23 | 24 | 23 | 23 | 26 | 26.0532 | 27.8 | 20.3 | 35 | 22 | 27 | 20 | 24.194 | 20.436 | 22.059
| Mean |  | 22.1062 | 24.6667 | 24.2333 | 34.3333 | 24.1667 | 23.6667 | 25.1667 | 24 | 23.6667 | 22.5 | 27.5 | 25.7356 | 28.1167 | 20.7833 | 32.3333 | 22 | 26.8 | 20.3333 | 24.7458 | 20.4732 | 22.459
| Median |  | 22.1497 | 25 | 24.7 | 34.5 | 24 | 24 | 25.5 | 24 | 24 | 22.5 | 27.5 | 25.72 | 28.1 | 20.8 | 32 | 22 | 26.65 | 20 | 24.7885 | 20.5016 | 22.406
| Std Dev. |  | 0.28731 | 0.8165 | 0.89368 | 0.8165 | 0.40825 | 0.5164 | 1.169 | 1.0954 | 0.5164 | 0.54772 | 1.0488 | 0.47634 | 0.7679 | 0.3656 | 1.8619 | 0.63246 | 0.28983 | 0.5164 | 0.36169 | 0.54999 | 0.40341
| Rel.Std.Dev. |  | 1.3% | 3.31% | 3.688% | 2.378% | 1.689% | 2.182% | 4.645% | 4.564% | 2.182% | 2.434% | 3.814% | 1.851% | 2.731% | 1.759% | 5.758% | 2.875% | 1.081% | 2.54% | 1.462% | 2.686% | 1.796%
| PDM3 |  | -7.305% | 3.431% | 1.614% | 43.96% | 1.334% | -0.7621% | 5.528% | 0.6356% | -0.7621% | -5.654% | 15.31% | 7.913% | 17.9% | -12.85% | 35.58% | -7.751% | 12.38% | -14.74% | 3.763% | -14.15% | -5.826%
| Z-Score (Absolute) |  | 0.81 | 0.2 | 0.03 | 4.01 | 0 | 0.2 | 0.39 | 0.07 | 0.2 | 0.66 | 1.32 | 0.62 | 1.56 | 1.34 | 3.22 | 0.86 | 1.04 | 1.51 | 0.23 | 1.46 | 0.67
| Analytical results for Cs in OREAS 605b (Certified Value 1.52 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 1.53 | 1.43 | 1.45 | 1.45 | 1.53 | 1.5 | 1.58 | 1.58 | 1.63 | 1.5338 | 1.46 | 1.4 | 1.65 | 1.44 | 1.413 | [8 empty] | 1
| 1 | 2 | 1.57 | 1.42 | 1.47 | 1.43 | 1.52 | 1.48 | 1.52 | 1.55 | 1.69 | 1.6607 | 1.54 | 1.4 | 1.62 | 1.39 | 1.437 | [8 empty] | 34
| 1 | 3 | 1.57 | 1.49 | 1.47 | 1.55 | 1.53 | 1.46 | 1.57 | 1.58 | 1.67 | 1.5551 | 1.36 | 1.42 | 1.66 | 1.44 | 1.443 | [8 empty] | 16
| 1 | 4 | 1.54 | 1.56 | 1.44 | 1.49 | 1.49 | 1.36 | 1.58 | 1.56 | 1.68 | 1.6353 | 1.49 | 1.42 | 1.65 | 1.43 | 1.461 | [8 empty] | 1.5171
| 1 | 5 | 1.58 | 1.59 | 1.45 | 1.35 | 1.55 | 1.51 | 1.49 | 1.56 | 1.66 | 1.5688 | 1.5 | 1.38 | 1.63 | 1.44 | 1.428 | [8 empty] | 81
| 1 | 6 | 1.59 | 1.55 | 1.39 | 1.6 | 1.5 | 1.48 | 1.5 | 1.57 | 1.59 | 1.5864 | 1.49 | 1.4 | 1.68 | 1.45 | 1.382
| Mean |  | 1.5633 | 1.5067 | 1.445 | 1.4783 | 1.52 | 1.465 | 1.54 | 1.5667 | 1.6533 | 1.59 | 1.4733 | 1.4033 | 1.6483 | 1.4317 | 1.4273
| Median |  | 1.57 | 1.52 | 1.45 | 1.47 | 1.525 | 1.48 | 1.545 | 1.565 | 1.665 | 1.5776 | 1.49 | 1.4 | 1.65 | 1.44 | 1.4325
| Std Dev. |  | 0.023381 | 0.071181 | 0.029496 | 0.089088 | 0.021909 | 0.054314 | 0.041473 | 0.012111 | 0.037238 | 0.048772 | 0.06121 | 0.015055 | 0.02137 | 0.02137 | 0.027325
| Rel.Std.Dev. |  | 1.496% | 4.724% | 2.041% | 6.026% | 1.441% | 3.707% | 2.693% | 0.773% | 2.252% | 3.067% | 4.155% | 1.073% | 1.296% | 1.493% | 1.914%
| PDM3 |  | 3.05% | -0.6856% | -4.75% | -2.553% | 0.1933% | -3.432% | 1.512% | 3.269% | 8.982% | 4.809% | -2.883% | -7.497% | 8.653% | -5.629% | -5.915%
| Z-Score (Absolute) |  | 0.64 | 0 | 0.69 | 0.32 | 0.15 | 0.47 | 0.37 | 0.67 | 1.65 | 0.94 | 0.37 | 1.16 | 1.59 | 0.84 | 0.89
| Analytical results for Cu in OREAS 605b (Certified Value 4.96 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | wt.%
|  |  | AR*OES | AR*OES | AR*OES | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*AAS | AR*AAS | AR*OES/MS | AR*OES/MS | AR*AAS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*OES | 2
|  |  | 0.5g | 01g | 01g | 0.2g | 0.5g | 0.5g | 0.4g | 0.5g | 0.5g | 01g | 10g | 01g | 01g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 10g | 0.5g | 0.2g | 0.5g | 0.25g | 3
| 1 | 1 | 4.8082 | > 1 | 4.821 | 4.59 | 4.84 | 5.07 | 5.07 | 4.99 | 5.05 | > 2 | > 2 | 4.965 | 4.766 | 5.011 | 5.6 | 4.91 | 4.883 | 5.11 | > 1 | 5.3132 | > 1 | 5.019 | 5.04 | 1
| 1 | 2 | 4.7382 | > 1 | 4.833 | 4.59 | 4.89 | 5.08 | NSS | 5.04 | 4.96 | > 2 | > 2 | 5.0244 | 5.163 | 4.977 | 5.7 | 4.8 | 4.833 | 5.09 | > 1 | 5.4232 | > 1 | 5.029 | 5.05 | 35
| 1 | 3 | 4.7574 | > 1 | 4.881 | 4.6 | 4.86 | 5.11 | NSS | 4.99 | 5.04 | > 2 | > 2 | 5 | 4.759 | 5.024 | 5.71 | 4.82 | 4.894 | 5.19 | > 1 | 5.0928 | > 1 | 5.006 | 4.99 | 16
| 1 | 4 | 4.7719 | > 1 | 4.89 | 4.6 | 4.98 | 5.11 | 4.95 | 5.01 | 5.05 | > 2 | > 2 | 5.0121 | 5.06 | 4.955 | 5.87 | 5.16 | 4.838 | 5.11 | > 1 | 5.291 | > 1 | 4.998 | 4.97 | 4.9609
| 1 | 5 | 4.7802 | > 1 | 4.833 | 4.58 | 4.88 | 5 | 4.96 | 4.97 | 4.94 | > 2 | > 2 | 4.9306 | 4.774 | 4.902 | 5.85 | 5.04 | 4.843 | 5.03 | > 1 | 5.325 | > 1 | 5.028 | 5.05 | 82
| 1 | 6 | 4.7406 | > 1 | 4.879 | 4.6 | 4.83 | 5.08 | 5.02 | 5.08 | 5.08 | > 2 | > 2 | 4.9203 | 5.12 | 4.915 | 5.86 | 4.81 | 4.965 | 5.09 | > 1 | 5.2561 | > 1 | 5.02 | 4.92
| Mean |  | 4.7661 |  | 4.8562 | 4.5933 | 4.88 | 5.075 | 5 | 5.0133 | 5.02 |  |  | 4.9754 | 4.9403 | 4.964 | 5.765 | 4.9233 | 4.876 | 5.1033 |  | 5.2835 |  | 5.0167 | 5.0033
| Median |  | 4.7647 |  | 4.856 | 4.595 | 4.87 | 5.08 | 4.99 | 5 | 5.045 |  |  | 4.9825 | 4.917 | 4.966 | 5.78 | 4.865 | 4.863 | 5.1 |  | 5.3021 |  | 5.0195 | 5.015
| Std Dev. |  | 0.026498 |  | 0.030308 | 0.008165 | 0.054037 | 0.040373 | 0.055976 | 0.040332 | 0.056214 |  |  | 0.043592 | 0.19345 | 0.049606 | 0.11113 | 0.14733 | 0.050351 | 0.05164 |  | 0.1089 |  | 0.012323 | 0.052789
| Rel.Std.Dev. |  | 0.556% |  | 0.6241% | 0.1778% | 1.107% | 0.7955% | 1.12% | 0.8045% | 1.12% |  |  | 0.8762% | 3.916% | 0.9993% | 1.928% | 2.993% | 1.033% | 1.012% |  | 2.061% |  | 0.2456% | 1.055%
| PDM3 |  | -3.926% |  | -2.111% | -7.409% | -1.63% | 2.301% | 0.7889% | 1.058% | 1.192% |  |  | 0.2928% | -0.4139% | 0.06317% | 16.21% | -0.7566% | -1.711% | 2.872% |  | 6.504% |  | 1.125% | 0.856%
| Z-Score (Absolute) |  | 1.97 | NA | 1.17 | 3.51 | 0.96 | 0.78 | 0.11 | 0.23 | 0.29 | NA | NA | 0.11 | 0.42 | 0.21 | 6.91 | 0.57 | 0.99 | 1.03 | NA | 2.63 | NA | 0.26 | 0.14
| Analytical results for Dy in OREAS 605b (Indicative Value 0.74 ppm) | [24 empty] | Indicative
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 10; 03: 23 | ppm
|  |  | 01: AR*OES/MS; 02: AR*OES/MS; 03: AR*MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.25g | 2
| 1 | 1 | 01: 0.69; 02: 0.79; 03: 0.646 | 1
| 1 | 2 | 01: 0.87; 02: 0.81; 03: 0.664 | 12
| 1 | 3 | 01: 0.63; 02: 0.8; 03: 0.675 | 16
| 1 | 4 | 01: 0.83; 02: 0.8; 03: 0.685 | 0.73539
| 1 | 5 | 01: 0.61; 02: 0.82; 03: 0.695 | 18
| 1 | 6 | 01: 0.79; 02: 0.8; 03: 0.632
| Mean |  | 01: 0.73667; 02: 0.80333; 03: 0.66617
| Median |  | 01: 0.74; 02: 0.8; 03: 0.6695
| Std Dev. |  | 01: 0.10857; 02: 0.010328; 03: 0.023845
| Rel.Std.Dev. |  | 01: 14.74%; 02: 1.286%; 03: 3.579%
| PDM3 |  | 01: 0.1738%; 02: 9.239%; 03: -9.413%
| Z-Score (Absolute) |  | 01: 0; 02: 0.67; 03: 0.71
| Analytical results for Er in OREAS 605b (Indicative Value 0.32 ppm) | [24 empty] | Indicative
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 10; 03: 23 | ppm
|  |  | 01: AR*OES/MS; 02: AR*OES/MS; 03: AR*MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.25g | 2
| 1 | 1 | 01: 0.33; 02: 0.34; 03: 0.306 | 1
| 1 | 2 | 01: 0.36; 02: 0.36; 03: 0.293 | 13
| 1 | 3 | 01: 0.28; 02: 0.37; 03: 0.309 | 16
| 1 | 4 | 01: 0.25; 02: 0.36; 03: 0.315 | 0.32094
| 1 | 5 | 01: 0.28; 02: 0.38; 03: 0.301 | 19
| 1 | 6 | 01: 0.29; 02: 0.36; 03: 0.293
| Mean |  | 01: 0.29833; 02: 0.36167; 03: 0.30283
| Median |  | 01: 0.285; 02: 0.36; 03: 0.3035
| Std Dev. |  | 01: 0.039707; 02: 0.013292; 03: 0.0088638
| Rel.Std.Dev. |  | 01: 13.31%; 02: 3.675%; 03: 2.927%
| PDM3 |  | 01: -7.045%; 02: 12.69%; 03: -5.643%
| Z-Score (Absolute) |  | 01: 0.67; 02: 8.82; 03: 0
| Analytical results for Eu in OREAS 605b (Indicative Value 0.16 ppm) | [24 empty] | Indicative
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 10; 03: 23 | ppm
|  |  | 01: AR*OES/MS; 02: AR*OES/MS; 03: AR*MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.25g | 2
| 1 | 1 | 01: 0.14; 02: 0.17; 03: 0.153 | 1
| 1 | 2 | 01: 0.15; 02: 0.18; 03: 0.158 | 14
| 1 | 3 | 01: 0.16; 02: 0.17; 03: 0.158 | 16
| 1 | 4 | 01: 0.19; 02: 0.18; 03: 0.165 | 0.16206
| 1 | 5 | 01: 0.15; 02: 0.18; 03: 0.157 | 20
| 1 | 6 | 01: 0.15; 02: 0.19; 03: 0.156
| Mean |  | 01: 0.15667; 02: 0.17833; 03: 0.15783
| Median |  | 01: 0.15; 02: 0.18; 03: 0.1575
| Std Dev. |  | 01: 0.017512; 02: 0.0075277; 03: 0.0039707
| Rel.Std.Dev. |  | 01: 11.18%; 02: 4.221%; 03: 2.516%
| PDM3 |  | 01: -3.325%; 02: 10.04%; 03: -2.605%
| Z-Score (Absolute) |  | 01: 0.67; 02: 11.85; 03: 0
| Analytical results for Fe in OREAS 605b (Certified Value 6.99 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | wt.%
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 3
| 1 | 1 | 5.7417 | 7.1 | 6.39 | 5.87 | 6.83 | 6.97 | 7.39 | 7.02 | 6.87 | 7.08 | 7.04 | 6.891 | 7.56 | 7.1 | 6.87 | 6.88 | 7.03 | 6.74 | 7.4217 | 6.5461 | 7.423 | 7.029 |  | 1
| 1 | 2 | 5.7897 | 6.91 | 6.42 | 5.84 | 6.77 | 6.85 | 7.32 | 6.87 | 6.86 | 7.16 | 7.18 | 7.0267 | 7.6 | 6.94 | 6.87 | 6.98 | 7.01 | 6.7 | 7.4148 | 6.5854 | 7.444 | 6.774 |  | #N/A
| 1 | 3 | 5.78 | 7.04 | 6.76 | 5.97 | 6.91 | 7.11 | 7.43 | 6.52 | 6.8 | 7.14 | 7.1 | 6.8189 | 7.46 | 6.96 | 6.89 | 6.92 | 6.94 | 6.55 | 7.3949 | 6.6705 | 7.46 | 6.978 |  | 16
| 1 | 4 | 5.7994 | 6.92 | 6.86 | 5.88 | 6.85 | 6.91 | 7.35 | 6.58 | 6.95 | 7.19 | 7.46 | 7.0972 | 7.7 | 7.04 | 7.43 | 6.93 | 7.01 | 6.63 | 7.5266 | 6.5205 | 7.435 | 7.199 |  | 6.9945
| 1 | 5 | 5.7919 | 6.96 | 6.94 | 5.95 | 6.85 | 6.86 | 7.34 | 6.15 | 6.93 | 7.08 | 7.19 | 6.8569 | 7.36 | 6.69 | 7.23 | 6.9 | 7.06 | 6.54 | 7.3139 | 6.6731 | 7.412 | 6.943 |  | 83
| 1 | 6 | 5.7617 | 6.86 | 6.72 | 5.82 | 6.66 | 6.95 | 6.78 | 6.4 | 6.85 | 7.11 | 6.89 | 7.0151 | 7.56 | 6.95 | 6.87 | 7.04 | 7.25 | 6.57 | 7.2248 | 6.6199 | 7.433 | 6.834
| Mean |  | 5.7774 | 6.965 | 6.6817 | 5.8883 | 6.8117 | 6.9417 | 7.2683 | 6.59 | 6.8767 | 7.1267 | 7.1433 | 6.951 | 7.54 | 6.9467 | 7.0267 | 6.9417 | 7.05 | 6.6217 | 7.3828 | 6.6026 | 7.4345 | 6.9595
| Median |  | 5.7849 | 6.94 | 6.74 | 5.875 | 6.84 | 6.93 | 7.345 | 6.55 | 6.865 | 7.125 | 7.14 | 6.953 | 7.56 | 6.955 | 6.88 | 6.925 | 7.02 | 6.6 | 7.4048 | 6.6026 | 7.434 | 6.9605
| Std Dev. |  | 0.021796 | 0.089387 | 0.2279 | 0.059805 | 0.086814 | 0.095167 | 0.24244 | 0.31559 | 0.055015 | 0.044572 | 0.19002 | 0.11057 | 0.1173 | 0.14024 | 0.24345 | 0.058793 | 0.10564 | 0.083287 | 0.10301 | 0.063419 | 0.016622 | 0.15031
| Rel.Std.Dev. |  | 0.3773% | 1.283% | 3.411% | 1.016% | 1.274% | 1.371% | 3.336% | 4.789% | 0.8% | 0.6254% | 2.66% | 1.591% | 1.556% | 2.019% | 3.465% | 0.847% | 1.498% | 1.258% | 1.395% | 0.9605% | 0.2236% | 2.16%
| PDM3 |  | -17.4% | -0.422% | -4.473% | -15.81% | -2.614% | -0.7556% | 3.915% | -5.783% | -1.685% | 1.889% | 2.128% | -0.6228% | 7.799% | -0.6841% | 0.4597% | -0.7556% | 0.7933% | -5.331% | 5.551% | -5.604% | 6.29% | -0.5006%
| Z-Score (Absolute) |  | 4.24 | 0.06 | 0.97 | 3.84 | 0.5 | 0.03 | 1.16 | 1.3 | 0.26 | 0.64 | 0.7 | 0.01 | 2.14 | 0.01 | 0.28 | 0.03 | 0.37 | 1.18 | 1.57 | 1.25 | 1.76 | 0.04
| Analytical results for Ga in OREAS 605b (Certified Value 8.65 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | 25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | 01g | [5 empty] | 3
| 1 | 1 | 9 | 8.1 | 11 | 8.56 | 8.22 | 9.64 | 8.6 | 8.27 | 8.55 | 9.14 | 9.4391 | 8.1 | 7.56 | 8 | 8.8 | 6.7 | 6.8582 | < 5 | [5 empty] | 1
| 1 | 2 | 8.7 | 7.9 | 11 | 8.79 | 8.04 | 9.14 | 8.49 | 8.33 | 8.51 | 9.23 | 10.5003 | 9.1 | 7.32 | 7.9 | 8.52 | 6.799 | 6.2376 | < 5 | [5 empty] | 39
| 1 | 3 | 8.9 | 8.6 | 11 | 8.7 | 8.63 | 9.27 | 9.38 | 8.57 | 8.5 | 9.25 | 9.6568 | 8.8 | 7.46 | 8.1 | 8.6 | 6.938 | 6.9649 | < 5 | [5 empty] | 16
| 1 | 4 | 8.7 | 8.6 | 11 | 8.49 | 8.63 | 8.88 | 8.72 | 8.56 | 8.58 | 9.44 | 10.3061 | 7.7 | 7.38 | 8.1 | 8.51 | 6.979 | 7.0463 | < 5 | [5 empty] | 8.6521
| 1 | 5 | 8.7 | 8.1 | 11 | 8.8 | 8.72 | 9.19 | 9.79 | 8.31 | 8.58 | 8.96 | 9.7364 | 8.9 | 7.15 | 7.9 | 8.44 | 6.887 | 6.6649 | < 5 | [5 empty] | 84
| 1 | 6 | 9.2 | 8.5 | 11 | 8.23 | 9.16 | 8.53 | 9.7 | 8.22 | 8.59 | 9 | 10.5289 | 8.5 | 7.31 | 7.9 | 8.67 | 6.639 | 7.003 | < 5
| Mean |  | 8.8667 | 8.3 | 11 | 8.595 | 8.5667 | 9.1083 | 9.1133 | 8.3767 | 8.5517 | 9.17 | 10.0279 | 8.5167 | 7.3633 | 7.9833 | 8.59 | 6.8237 | 6.7958
| Median |  | 8.8 | 8.3 | 11 | 8.63 | 8.63 | 9.165 | 9.05 | 8.32 | 8.565 | 9.185 | 10.0212 | 8.65 | 7.35 | 7.95 | 8.56 | 6.843 | 6.9115
| Std Dev. |  | 0.20656 | 0.30332 | 0 | 0.21751 | 0.39526 | 0.37499 | 0.57964 | 0.15069 | 0.038687 | 0.17709 | 0.47349 | 0.53072 | 0.14038 | 0.098319 | 0.12992 | 0.13507 | 0.30557
| Rel.Std.Dev. |  | 2.33% | 3.654% | 0% | 2.531% | 4.614% | 4.117% | 6.36% | 1.799% | 0.4524% | 1.931% | 4.722% | 6.232% | 1.906% | 1.232% | 1.512% | 1.979% | 4.497%
| PDM3 |  | 2.48% | -4.07% | 27.14% | -0.6601% | -0.9876% | 5.273% | 5.331% | -3.184% | -1.161% | 5.986% | 15.9% | -1.565% | -14.9% | -7.73% | -0.7179% | -21.13% | -21.45%
| Z-Score (Absolute) |  | 0.38 | 0.3 | NA | 0.05 | 0.02 | 0.67 | 0.68 | 0.21 | 0 | 0.75 | 1.79 | 0.04 | 1.44 | 0.69 | 0.05 | 2.09 | 2.13 | 7.33
Results from laboratory 4 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 605b (Indicative Value 0.94 ppm) | [24 empty] | Indicative
Gd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab | [5 empty] | 1
| No | No | 02: 03; 03: 10; 04: 23 | [5 empty] | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS | [5 empty] | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g | [5 empty] | 2
| 1 | 1 | 02: 0.97; 03: 1.16; 04: 0.948 | [5 empty] | 1
| 1 | 2 | 02: 1; 03: 1.19; 04: 0.947 | [5 empty] | 15
| 1 | 3 | 02: 0.94; 03: 1.16; 04: 0.908 | [5 empty] | 16
| 1 | 4 | 02: 0.86; 03: 1.11; 04: 0.974 | [5 empty] | 0.94425
| 1 | 5 | 02: 0.93; 03: 1.14; 04: 0.941 | [5 empty] | 21
| 1 | 6 | 02: 0.96; 03: 1.19; 04: 0.953
| Mean |  | 02: 0.94333; 03: 1.1583; 04: 0.94517
| Median |  | 02: 0.95; 03: 1.16; 04: 0.9475
| Std Dev. |  | 02: 0.04761; 03: 0.030605; 04: 0.021461
| Rel.Std.Dev. |  | 02: 5.047%; 03: 2.642%; 04: 2.271%
| PDM3 |  | 02: -0.09708%; 03: 22.67%; 04: 0.09708%
| Z-Score (Absolute) |  | 02: 0.67; 03: 78.4; 04: 0
| Analytical results for Ge in OREAS 605b (Indicative Value 0.6 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 03 | 05 | 06 | 07 | 08 | 09 | 14 | 15 | 20 | 21 | 23 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [12 empty] | 2
| 1 | 1 | 1.3 | 0.52 | 0.56 | 0.93 | 0.44 | 0.49 | 0.1 | 0.3 | 1.1 | 1.02 | 0.305 | [12 empty] | 1
| 1 | 2 | 1.3 | 0.59 | 0.47 | 0.84 | 0.42 | 0.44 | 0.1 | 0.34 | 1 | 1.05 | 0.265 | [12 empty] | 16
| 1 | 3 | 1.1 | 0.58 | 0.47 | 1 | 0.62 | 0.45 | 0.1 | 0.33 | 1 | 0.94 | 0.286 | [12 empty] | 16
| 1 | 4 | 1.5 | 0.51 | 0.75 | 0.9 | 0.55 | 0.41 | 0.2 | 0.28 | 1 | 0.98 | 0.268 | [12 empty] | 0.5981
| 1 | 5 | 1.3 | 0.52 | 0.75 | 1.22 | 0.63 | 0.4 | 0.2 | 0.37 | 1.1 | 1.02 | 0.279 | [12 empty] | 22
| 1 | 6 | 1.3 | 0.63 | 0.8 | 1.08 | 0.64 | 0.37 | 0.2 | 0.33 | 1.2 | 1.02 | 0.268
| Mean |  | 1.3 | 0.55833 | 0.63333 | 0.995 | 0.55 | 0.42667 | 0.15 | 0.325 | 1.0667 | 1.005 | 0.2785
| Median |  | 1.3 | 0.55 | 0.655 | 0.965 | 0.585 | 0.425 | 0.15 | 0.33 | 1.05 | 1.02 | 0.2735
| Std Dev. |  | 0.12649 | 0.048751 | 0.15082 | 0.13795 | 0.098387 | 0.042269 | 0.054772 | 0.031464 | 0.08165 | 0.038859 | 0.015241
| Rel.Std.Dev. |  | 9.73% | 8.732% | 23.81% | 13.86% | 17.89% | 9.907% | 36.51% | 9.681% | 7.655% | 3.867% | 5.473%
| PDM3 |  | 117.4% | -6.649% | 5.89% | 66.36% | -8.043% | -28.66% | -74.92% | -45.66% | 78.34% | 68.03% | -53.44%
| Z-Score (Absolute) |  | NA | 0.02 | 0.3 | 1.67 | 0.02 | 0.48 | NA | 0.87 | NA | 1.7 | 1.04
Results from laboratories 3, 14 & 20 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 605b (Certified Value 0.44 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [8 empty] | 2
| 1 | 1 | 0.53 | 0.39 | 0.45 | 0.43 | 0.44 | 0.45 | 0.44 | 0.53 | 0.47 | 0.48079 | 0.43 | 0.37 | 0.39 | 0.412 | 0.4 | [8 empty] | 1
| 1 | 2 | 0.51 | 0.42 | 0.46 | 0.39 | 0.44 | 0.46 | 0.44 | 0.51 | 0.52 | 0.4803 | 0.37 | 0.38 | 0.38 | 0.431 | 0.431 | [8 empty] | 25
| 1 | 3 | 0.53 | 0.44 | 0.44 | 0.4 | 0.45 | 0.42 | 0.45 | 0.53 | 0.49 | 0.47172 | 0.38 | 0.37 | 0.37 | 0.435 | 0.415 | [8 empty] | 16
| 1 | 4 | 0.52 | 0.35 | 0.45 | 0.5 | 0.44 | 0.39 | 0.45 | 0.55 | 0.49 | 0.48918 | 0.39 | 0.35 | 0.39 | 0.446 | 0.399 | [8 empty] | 0.43866
| 1 | 5 | 0.51 | 0.33 | 0.45 | 0.4 | 0.47 | 0.46 | 0.43 | 0.52 | 0.49 | 0.47341 | 0.4 | 0.34 | 0.4 | 0.414 | 0.383 | [8 empty] | 85
| 1 | 6 | 0.55 | 0.49 | 0.39 | 0.37 | 0.4 | 0.46 | 0.44 | 0.55 | 0.47 | 0.48936 | 0.4 | 0.36 | 0.38 | 0.419 | 0.392
| Mean |  | 0.525 | 0.40333 | 0.44 | 0.415 | 0.44 | 0.44 | 0.44167 | 0.53167 | 0.48833 | 0.48079 | 0.395 | 0.36167 | 0.385 | 0.42617 | 0.40333
| Median |  | 0.525 | 0.405 | 0.45 | 0.4 | 0.44 | 0.455 | 0.44 | 0.53 | 0.49 | 0.48055 | 0.395 | 0.365 | 0.385 | 0.425 | 0.3995
| Std Dev. |  | 0.015166 | 0.059217 | 0.025298 | 0.045935 | 0.022804 | 0.028983 | 0.0075277 | 0.016021 | 0.018348 | 0.0074937 | 0.020736 | 0.01472 | 0.010488 | 0.013378 | 0.017166
| Rel.Std.Dev. |  | 2.889% | 14.68% | 5.75% | 11.07% | 5.183% | 6.587% | 1.704% | 3.013% | 3.757% | 1.559% | 5.25% | 4.07% | 2.724% | 3.139% | 4.256%
| PDM3 |  | 19.68% | -8.054% | 0.3045% | -5.395% | 0.3045% | 0.3045% | 0.6845% | 21.2% | 11.32% | 9.604% | -9.954% | -17.55% | -12.23% | -2.849% | -8.054%
| Z-Score (Absolute) |  | 1.56 | 0.67 | 0 | 0.46 | 0 | 0 | 0.03 | 1.69 | 0.89 | 0.75 | 0.83 | 1.44 | 1.01 | 0.25 | 0.67
| Analytical results for Hg in OREAS 605b (Certified Value 2.2 ppm) | [24 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 14 | 15 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 10g | 0.25g | 0.2g | 01g | 0.25g |  |  |  |  | 3
| 1 | 1 | 2.23 | 2.137 | < 5 | 2.07 | 1.89 | 2.03 | 2.21 | 2.04 | 2.3 | 1.8873 | 1.312 | 2.44 | 2.46 | 2.01 | < 5 | 2.358 | < 2 | < 5 | 2.987 |  |  |  |  | 1
| 1 | 2 | 2.23 | 1.908 | < 5 | 2.01 | 2.14 | 2.02 | 2.24 | 2.1 | 2.5 | 1.9482 | 1.338 | 2.37 | 2.47 | 2.14 | < 5 | 2.424 | < 2 | < 5 | 2.725 |  |  |  |  | 7
| 1 | 3 | 2.31 | 2.124 | < 5 | 2.11 | 2.19 | 2.32 | 2.08 | 2.04 | 2.4 | 1.9215 | 1.346 | 2.35 | 2.5 | 2.16 | < 5 | 2.436 | < 2 | < 5 | 3.032 |  |  |  |  | 16
| 1 | 4 | 2.25 | 2.16 | < 5 | 2.1 | 2.2 | 2.13 | 1.95 | 2.06 | 2.3 | 1.9279 | 1.301 | 2.37 | 2.51 | 2.11 | < 5 | 2.353 | < 2 | < 5 | 3.171 |  |  |  |  | 2.1979
| 1 | 5 | 2.22 | 2.226 | < 5 | 2.04 | 2.06 | 2.13 | 2.13 | 2.03 | 2.4 | 1.8956 | 1.325 | 2.35 | 2.41 | 1.97 | < 5 | 2.478 | < 2 | < 5 | 2.9 |  |  |  |  | 86
| 1 | 6 | 2.22 | 2.131 | < 5 | 2.07 | 2.23 | 2.24 | 2.14 | 1.97 | 2.4 | 1.9238 | 1.438 | 2.47 | 2.36 | 2.11 | < 5 | 2.387 | < 2 | < 5 | 2.793
| Mean |  | 2.2433 | 2.1143 |  | 2.0667 | 2.1183 | 2.145 | 2.125 | 2.04 | 2.3833 | 1.9174 | 1.3433 | 2.3917 | 2.4517 | 2.0833 |  | 2.406 |  |  | 2.9347
| Median |  | 2.23 | 2.134 |  | 2.07 | 2.165 | 2.13 | 2.135 | 2.04 | 2.4 | 1.9227 | 1.3315 | 2.37 | 2.465 | 2.11 |  | 2.4055 |  |  | 2.9435
| Std Dev. |  | 0.034448 | 0.10772 |  | 0.037238 | 0.12671 | 0.11743 | 0.1033 | 0.042426 | 0.075277 | 0.022349 | 0.049208 | 0.050761 | 0.057067 | 0.075807 |  | 0.048748 |  |  | 0.16328
| Rel.Std.Dev. |  | 1.536% | 5.095% |  | 1.802% | 5.982% | 5.475% | 4.861% | 2.08% | 3.158% | 1.166% | 3.663% | 2.122% | 2.328% | 3.639% |  | 2.026% |  |  | 5.564%
| PDM3 |  | 2.066% | -3.803% |  | -5.972% | -3.621% | -2.408% | -3.318% | -7.185% | 8.436% | -12.76% | -38.88% | 8.815% | 11.54% | -5.214% |  | 9.467% |  |  | 33.52%
| Z-Score (Absolute) |  | 0.28 | 0.09 | 1 | 0.22 | 0.08 | 0 | 0.06 | 0.3 | 0.67 | 0.64 | 2.27 | 0.7 | 0.87 | 0.17 | 1 | 0.74 | 3.24 | 1 | 2.23
| Analytical results for Ho in OREAS 605b (Indicative Value 0.13 ppm) | [24 empty] | Indicative
Ho | Round | Replicate | 02: Lab; 03: Lab; 04: Lab |  |  |  |  | 1
| No | No | 02: 03; 03: 10; 04: 23 |  |  |  |  | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS |  |  |  |  | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g |  |  |  |  | 2
| 1 | 1 | 02: 0.11; 03: 0.14; 04: 0.116 |  |  |  |  | 1
| 1 | 2 | 02: 0.13; 03: 0.14; 04: 0.114 |  |  |  |  | 10
| 1 | 3 | 02: 0.13; 03: 0.14; 04: 0.115 |  |  |  |  | 16
| 1 | 4 | 02: 0.14; 03: 0.14; 04: 0.119 |  |  |  |  | 0.12817
| 1 | 5 | 02: 0.08; 03: 0.14; 04: 0.116 |  |  |  |  | 16
| 1 | 6 | 02: 0.13; 03: 0.14; 04: 0.119
| Mean |  | 02: 0.12; 03: 0.14; 04: 0.1165
| Median |  | 02: 0.13; 03: 0.14; 04: 0.116
| Std Dev. |  | 02: 0.021909; 03: 0; 04: 0.0020736
| Rel.Std.Dev. |  | 02: 18.26%; 03: 0%; 04: 1.78%
| PDM3 |  | 02: -6.372%; 03: 9.233%; 04: -9.103%
| Z-Score (Absolute) |  | 02: 0; 03: 3.85; 04: 0.67
| Analytical results for In in OREAS 605b (Certified Value 3.8 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [7 empty] | 3
| 1 | 1 | 3.23 | 3.64 | < 5 | 3.78 | 3.6 | 3.76 | 3.97 | 4.08 | 3.8 | 3.94 | 3.7994 | 3.45 | 3.64 | 3.93 | 3.63 | 3.478 | [7 empty] | 1
| 1 | 2 | 3.29 | 3.77 | < 5 | 3.91 | 3.71 | 3.53 | 3.86 | 4.07 | 3.88 | 4 | 3.8467 | 3.24 | 3.54 | 3.89 | 3.63 | 3.656 | [7 empty] | 27
| 1 | 3 | 3.28 | 4.04 | < 5 | 3.89 | 3.78 | 3.92 | 3.81 | 4.1 | 3.89 | 4.02 | 3.7423 | 3.38 | 3.61 | 3.96 | 3.63 | 3.633 | [7 empty] | 16
| 1 | 4 | 3.33 | 4.1 | < 5 | 3.81 | 3.67 | 3.67 | 3.56 | 4.12 | 3.82 | 4.24 | 3.893 | 3.23 | 3.62 | 3.94 | 3.66 | 3.634 | [7 empty] | 3.8013
| 1 | 5 | 3.33 | 3.78 | < 5 | 3.83 | 3.58 | 3.72 | 3.95 | 3.85 | 3.83 | 4.02 | 3.722 | 3.32 | 3.58 | 3.82 | 3.66 | 3.57 | [7 empty] | 87
| 1 | 6 | 3.4 | 4.01 | < 5 | 3.77 | 3.89 | 3.66 | 3.88 | 3.87 | 3.75 | 3.87 | 3.7929 | 3.56 | 3.59 | 3.88 | 3.69 | 3.498
| Mean |  | 3.31 | 3.89 |  | 3.8317 | 3.705 | 3.71 | 3.8383 | 4.015 | 3.8283 | 4.015 | 3.7994 | 3.3633 | 3.5967 | 3.9033 | 3.65 | 3.5782
| Median |  | 3.31 | 3.895 |  | 3.82 | 3.69 | 3.695 | 3.87 | 4.075 | 3.825 | 4.01 | 3.7961 | 3.35 | 3.6 | 3.91 | 3.645 | 3.6015
| Std Dev. |  | 0.057619 | 0.18439 |  | 0.057417 | 0.1164 | 0.129 | 0.14851 | 0.12145 | 0.051929 | 0.12454 | 0.063722 | 0.12754 | 0.035024 | 0.050859 | 0.024495 | 0.075774
| Rel.Std.Dev. |  | 1.741% | 4.74% |  | 1.498% | 3.142% | 3.477% | 3.869% | 3.025% | 1.356% | 3.102% | 1.677% | 3.792% | 0.9738% | 1.303% | 0.6711% | 2.118%
| PDM3 |  | -12.92% | 2.334% |  | 0.7996% | -2.533% | -2.401% | 0.9749% | 5.623% | 0.7119% | 5.623% | -0.04967% | -11.52% | -5.383% | 2.685% | -3.98% | -5.869%
| Z-Score (Absolute) |  | 2.11 | 0.64 | 5.96 | 0.37 | 0.24 | 0.21 | 0.4 | 1.24 | 0.35 | 1.24 | 0.21 | 1.86 | 0.75 | 0.71 | 0.5 | 0.84
| Analytical results for Ir in OREAS 605b (Indicative Value < 5 ppb) | [24 empty] | Indicative
Ir | Round | Replicate | 02: Lab | [7 empty] | 1
| No | No | 02: 23 | [7 empty] | ppb
|  |  | 02: AR*MS | [7 empty] | 2
|  |  | 02: 0.25g | [7 empty] | 2
| 1 | 1 | 02: < 5 | [7 empty] | 1
| 1 | 2 | 02: < 5 | [7 empty] | 8
| 1 | 3 | 02: < 5 | [7 empty] | 16
| 1 | 4 | 02: < 5 | [7 empty] | < 5
| 1 | 5 | 02: < 5 | [7 empty] | 17
| 1 | 6 | 02: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for K in OREAS 605b (Certified Value 0.247 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.2989 | 0.25 | 0.22 | 0.25 | 0.25 | 0.25 | 0.23 | 0.23 | 0.264 | 0.2486 | 0.24252 | 0.24 | 0.22 | 0.18 | 0.25 | 0.273 | 0.25 | 0.23404 | 0.23491 | 0.2535 | 0.24 |  |  | 1
| 1 | 2 | 0.2912 | 0.24 | 0.22 | 0.25 | 0.25 | 0.25 | 0.23 | 0.23 | 0.257 | 0.2528 | 0.2518 | 0.25 | 0.21 | 0.18 | 0.25 | 0.272 | 0.25 | 0.23593 | 0.23972 | 0.2635 | 0.247 |  |  | #N/A
| 1 | 3 | 0.2977 | 0.25 | 0.22 | 0.25 | 0.25 | 0.25 | 0.21 | 0.23 | 0.264 | 0.2545 | 0.2405 | 0.24 | 0.21 | 0.18 | 0.24 | 0.278 | 0.25 | 0.23682 | 0.24391 | 0.2625 | 0.246 |  |  | 16
| 1 | 4 | 0.2906 | 0.25 | 0.22 | 0.25 | 0.24 | 0.25 | 0.21 | 0.23 | 0.265 | 0.2551 | 0.25461 | 0.25 | 0.21 | 0.19 | 0.24 | 0.274 | 0.25 | 0.2419 | 0.23973 | 0.2651 | 0.243 |  |  | 0.24659
| 1 | 5 | 0.2854 | 0.25 | 0.21 | 0.25 | 0.24 | 0.25 | 0.2 | 0.23 | 0.26 | 0.2443 | 0.24684 | 0.24 | 0.2 | 0.18 | 0.26 | 0.268 | 0.24 | 0.23841 | 0.2384 | 0.2527 | 0.245 |  |  | 88
| 1 | 6 | 0.2921 | 0.25 | 0.22 | 0.24 | 0.24 | 0.22 | 0.21 | 0.23 | 0.261 | 0.2455 | 0.25046 | 0.24 | 0.21 | 0.18 | 0.24 | 0.277 | 0.25 | 0.22826 | 0.2402 | 0.2579 | 0.246
| Mean |  | 0.29265 | 0.24833 | 0.21833 | 0.24833 | 0.245 | 0.245 | 0.215 | 0.23 | 0.26183 | 0.25013 | 0.24779 | 0.24333 | 0.21 | 0.18167 | 0.24667 | 0.27367 | 0.24833 | 0.23589 | 0.23948 | 0.2592 | 0.2445
| Median |  | 0.29165 | 0.25 | 0.22 | 0.25 | 0.245 | 0.25 | 0.21 | 0.23 | 0.2625 | 0.2507 | 0.24865 | 0.24 | 0.21 | 0.18 | 0.245 | 0.2735 | 0.25 | 0.23638 | 0.23972 | 0.2602 | 0.2455
| Std Dev. |  | 0.0049746 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0054772 | 0.012247 | 0.012247 | 0 | 0.0030605 | 0.0046625 | 0.0055037 | 0.005164 | 0.0063246 | 0.0040825 | 0.008165 | 0.0036148 | 0.0040825 | 0.0045842 | 0.0029098 | 0.0053032 | 0.0025884
| Rel.Std.Dev. |  | 1.7% | 1.644% | 1.87% | 1.644% | 2.236% | 4.999% | 5.696% | 0% | 1.169% | 1.864% | 2.221% | 2.122% | 3.012% | 2.247% | 3.31% | 1.321% | 1.644% | 1.943% | 1.215% | 2.046% | 1.059%
| PDM3 |  | 18.68% | 0.7075% | -11.46% | 0.7075% | -0.6442% | -0.6442% | -12.81% | -6.727% | 6.182% | 1.438% | 0.487% | -1.32% | -14.84% | -26.33% | 0.03165% | 10.98% | 0.7075% | -4.337% | -2.883% | 5.114% | -0.847%
| Z-Score (Absolute) |  | 5.82 | 0.41 | 3.26 | 0.41 | 0 | 0 | 3.66 | 1.83 | 2.06 | 0.63 | 0.34 | 0.2 | 4.28 | 7.74 | 0.2 | 3.5 | 0.41 | 1.11 | 0.67 | 1.73 | 0.06
| Analytical results for La in OREAS 605b (Indicative Value 3.58 ppm) | [24 empty] | Indicative
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 2
| 1 | 1 | < 10 | 7.04 | 2.8 | < 5 | 3.3 | 3.1 | 3.8 | 2.8 | 2.8 | 6.39 | 2.754 | 5.7459 | 3.4 | 6.1 | 2.2 | 4.3 | 4.18 | 3.3 | 6.671 | 6.1347 | 6.673 | 2.77 |  | 1
| 1 | 2 | < 10 | 7.4 | 2.7 | < 5 | 3.2 | 3.2 | 3.1 | 2.8 | 2.9 | 6.48 | 2.855 | 5.5669 | 3.5 | 6.1 | 2.3 | 4.1 | 3.85 | 3.1 | 6.864 | 6.4883 | 7.021 | 2.367 |  | 5
| 1 | 3 | < 10 | 7.4 | 3 | < 5 | 3.4 | 3.3 | 3.2 | 3 | 3.2 | 6.52 | 2.668 | 6.0167 | 3.5 | 6.2 | 1.9 | 4.1 | 3.6 | 3.1 | 6.561 | 6.4131 | 6.896 | 2.788 |  | 16
| 1 | 4 | < 10 | 6.91 | 3 | < 5 | 3.2 | 3.4 | 3.4 | 2.8 | 3.4 | 6.52 | 2.94 | 6.2205 | 3.4 | 6.1 | 1.9 | 4.2 | 4.64 | 3.1 | 6.664 | 6.4845 | 7.042 | 3.185 |  | 3.5755
| 1 | 5 | < 10 | 6.98 | 2.9 | < 5 | 3.4 | 3.2 | 3.4 | 2.9 | 2.7 | 6.48 | 2.482 | 6.3389 | 3.7 | 5.8 | 2.2 | 4.2 | 4.91 | 3.1 | 6.594 | 6.1186 | 6.978 | 3.313 |  | 18
| 1 | 6 | < 10 | 7.37 | 3 | < 5 | 3.1 | 3.4 | 3.2 | 2.9 | 3 | 6.56 | 2.755 | 6.193 | 3.5 | 6.1 | 2.5 | 4.3 | 4.09 | 2.9 | 6.601 | 6.1423 | 6.638 | 3.147
| Mean |  |  | 7.1833 | 2.9 |  | 3.2667 | 3.2667 | 3.35 | 2.8667 | 3 | 6.4917 | 2.7423 | 6.0136 | 3.5 | 6.0667 | 2.1667 | 4.2 | 4.2117 | 3.1 | 6.6592 | 6.2969 | 6.8747 | 2.9283
| Median |  |  | 7.205 | 2.95 |  | 3.25 | 3.25 | 3.3 | 2.85 | 2.95 | 6.5 | 2.7545 | 6.1048 | 3.5 | 6.1 | 2.2 | 4.2 | 4.135 | 3.1 | 6.6325 | 6.2777 | 6.937 | 2.9675
| Std Dev. |  |  | 0.23036 | 0.12649 |  | 0.12111 | 0.12111 | 0.251 | 0.08165 | 0.26077 | 0.058109 | 0.15834 | 0.30069 | 0.10954 | 0.13663 | 0.23381 | 0.089443 | 0.48824 | 0.12649 | 0.10898 | 0.18292 | 0.17733 | 0.35257
| Rel.Std.Dev. |  |  | 3.207% | 4.362% |  | 3.707% | 3.707% | 7.492% | 2.848% | 8.692% | 0.8951% | 5.774% | 5% | 3.13% | 2.252% | 10.79% | 2.13% | 11.59% | 4.08% | 1.637% | 2.905% | 2.579% | 12.04%
| PDM3 |  |  | 100.9% | -18.89% |  | -8.638% | -8.638% | -6.307% | -19.82% | -16.1% | 81.56% | -23.3% | 68.19% | -2.112% | 69.67% | -39.4% | 17.47% | 17.79% | -13.3% | 86.24% | 76.11% | 92.27% | -18.1%
| Z-Score (Absolute) |  | 1.36 | 3.25 | 0.45 | 0.8 | 0.14 | 0.14 | 0.06 | 0.48 | 0.37 | 2.65 | 0.59 | 2.24 | 0.06 | 2.28 | 1.09 | 0.67 | 0.68 | 0.28 | 2.79 | 2.48 | 2.98 | 0.43
| Analytical results for Li in OREAS 605b (Certified Value 8.63 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 26 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 0.25g |  |  | 2
| 1 | 1 | < 10 | 8.7 | 8.9 | 9 | 9.2 | 8.5 | 8.5 | 8.5 | 8.3 | 10.1 | 8.6 | 10.3612 | 9.7 | 7.5 | 11 | 10 | 9.03 | 9 | 7.071 | 8.1591 | 6.608 |  |  | 1
| 1 | 2 | < 10 | 8.3 | 8.6 | 9 | 9.5 | 8.7 | 8.2 | 8.3 | 7.5 | 10 | 8.8 | 10.2732 | 10 | 7.3 | 10 | 10 | 9.34 | 9 | 7.119 | 8.2934 | 7.065 |  |  | #N/A
| 1 | 3 | < 10 | 8.1 | 9.4 | 8 | 9.6 | 8.9 | 9 | 8.2 | 7.9 | 10.1 | 8.6 | 10.0589 | 10 | 7.3 | 11 | 9 | 9.57 | 9 | 6.931 | 8.3978 | 6.588 |  |  | 16
| 1 | 4 | < 10 | 7.9 | 9.2 | 9 | 9 | 8.5 | 8.2 | 8.1 | 8.1 | 10.1 | 9.5 | 10.4936 | 10.1 | 7.3 | 11 | 9 | 9.47 | 9 | 7.08 | 8.2869 | 7.589 |  |  | 8.6269
| 1 | 5 | < 10 | 7.9 | 8.6 | 8 | 9.7 | 8.5 | 8.5 | 9 | 7.8 | 10.3 | 8.4 | 10.0967 | 9.8 | 7 | 11 | 10 | 9.26 | 9 | 7.024 | 8.4374 | 6.75 |  |  | 89
| 1 | 6 | < 10 | 8.6 | 9.2 | 9 | 9.1 | 8.8 | 8 | 9.5 | 7.6 | 10.1 | 8.9 | 10.3555 | 9.9 | 7.3 | 11 | 9 | 9.62 | 9 | 6.764 | 8.2657 | 6.744
| Mean |  |  | 8.25 | 8.9833 | 8.6667 | 9.35 | 8.65 | 8.4 | 8.6 | 7.8667 | 10.1167 | 8.8 | 10.2732 | 9.9167 | 7.2833 | 10.8333 | 9.5 | 9.3817 | 9 | 6.9982 | 8.3067 | 6.8907
| Median |  |  | 8.2 | 9.05 | 9 | 9.35 | 8.6 | 8.35 | 8.4 | 7.85 | 10.1 | 8.7 | 10.3144 | 9.95 | 7.3 | 11 | 9.5 | 9.405 | 9 | 7.0475 | 8.2902 | 6.747
| Std Dev. |  |  | 0.34496 | 0.33714 | 0.5164 | 0.2881 | 0.17607 | 0.35214 | 0.54406 | 0.30111 | 0.098319 | 0.38471 | 0.16744 | 0.1472 | 0.16021 | 0.40825 | 0.54772 | 0.21922 | 0 | 0.13162 | 0.099428 | 0.38233
| Rel.Std.Dev. |  |  | 4.181% | 3.753% | 5.958% | 3.081% | 2.035% | 4.192% | 6.326% | 3.828% | 0.9719% | 4.372% | 1.63% | 1.484% | 2.2% | 3.768% | 5.766% | 2.337% | 0% | 1.881% | 1.197% | 5.548%
| PDM3 |  |  | -4.369% | 4.131% | 0.4606% | 8.381% | 0.2674% | -2.631% | -0.3122% | -8.813% | 17.27% | 2.006% | 19.08% | 14.95% | -15.57% | 25.58% | 10.12% | 8.749% | 4.324% | -18.88% | -3.712% | -20.13%
| Z-Score (Absolute) |  | 3.24 | 0.31 | 0.34 | NA | 0.67 | 0.04 | 0.18 | 0 | 0.66 | 1.36 | 0.18 | 1.5 | 1.18 | 1.18 | NA | NA | 0.7 | NA | 1.44 | 0.26 | 1.54
Results from laboratories 4, 19, 20 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 605b (Certified Value 36.1 ppb) | [24 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  |  | 1
| No | No | 01: 03; 02: 10; 03: 20; 04: 21; 05: 23 |  |  | ppb
|  |  | 01: AR*OES/MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS |  |  | 1
|  |  | 01: 0.5g; 02: 01g; 03: 0.25g; 04: 0.25g; 05: 0.25g |  |  | 1
| 1 | 1 | 01: 40; 02: 40; 03: 40; 04: 36.4; 05: 23 |  |  | 1
| 1 | 2 | 01: 50; 02: 40; 03: 40; 04: 37.3; 05: 36 |  |  | 29
| 1 | 3 | 01: 30; 02: 40; 03: 30; 04: 40.1; 05: 31 |  |  | 16
| 1 | 4 | 01: 30; 02: 40; 03: 40; 04: 35.7; 05: 40 |  |  | 36.0987
| 1 | 5 | 01: 30; 02: 40; 03: 30; 04: 37.4; 05: 26 |  |  | 90
| 1 | 6 | 01: 30; 02: 40; 03: 40; 04: 36.5; 05: 37
| Mean |  | 01: 35; 02: 40; 03: 36.6667; 04: 37.2333; 05: 32.1667
| Median |  | 01: 30; 02: 40; 03: 40; 04: 36.9; 05: 33.5
| Std Dev. |  | 01: 8.3666; 02: 0; 03: 5.164; 04: 1.5384; 05: 6.6758
| Rel.Std.Dev. |  | 01: 23.9%; 02: 0%; 03: 14.08%; 04: 4.132%; 05: 20.75%
| PDM3 |  | 01: -3.044%; 02: 10.81%; 03: 1.573%; 04: 3.143%; 05: -10.89%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.35; 03: 0; 04: 0.23; 05: 1.82
| Analytical results for Mg in OREAS 605b (Certified Value 1322 ppm) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 0
| 1 | 1 | 1400 | 1400 | 1200 | 1200 | 1200 | 1300 | 1200 | 1200 | 1600 | 1300 | 1111.4535 | 1400 | 1300 | 1400 | 1400 | 1450 | 1300 | 1856.36 | 1307.1178 | 1450 | 1330 |  |  | 1
| 1 | 2 | 1400 | 1400 | 1200 | 1200 | 1200 | 1300 | 1200 | 1200 | 1500 | 1300 | 1208.8305 | 1500 | 1300 | 1400 | 1400 | 1450 | 1300 | 1872.431 | 1309.69 | 1416 | 1350 |  |  | #N/A
| 1 | 3 | 1400 | 1500 | 1200 | 1300 | 1200 | 1300 | 1100 | 1200 | 1500 | 1300 | 1147.1955 | 1500 | 1300 | 1300 | 1400 | 1460 | 1300 | 1777.192 | 1322.1653 | 1418 | 1340 |  |  | 16
| 1 | 4 | 1400 | 1400 | 1200 | 1300 | 1200 | 1300 | 1100 | 1200 | 1500 | 1500 | 1199.216 | 1500 | 1300 | 1500 | 1300 | 1460 | 1400 | 1858.171 | 1295.7833 | 1458 | 1320 |  |  | 1321.5401
| 1 | 5 | 1300 | 1500 | 1100 | 1200 | 1200 | 1300 | 1100 | 1200 | 1500 | 1300 | 1143.7014 | 1400 | 1200 | 1400 | 1400 | 1470 | 1300 | 1789.542 | 1315.1277 | 1422 | 1320 |  |  | 91
| 1 | 6 | 1400 | 1400 | 1100 | 1200 | 1200 | 1200 | 1100 | 1200 | 1500 | 1300 | 1190.1611 | 1500 | 1300 | 1400 | 1400 | 1510 | 1300 | 1766.628 | 1311.3691 | 1411 | 1340
| Mean |  | 1383.3333 | 1433.3333 | 1166.6667 | 1233.3333 | 1200 | 1283.3333 | 1133.3333 | 1200 | 1516.6667 | 1333.3333 | 1166.7597 | 1466.6667 | 1283.3333 | 1400 | 1383.3333 | 1466.6667 | 1316.6667 | 1820.054 | 1310.2088 | 1429.1667 | 1333.3333
| Median |  | 1400 | 1400 | 1200 | 1200 | 1200 | 1300 | 1100 | 1200 | 1500 | 1300 | 1168.6783 | 1500 | 1300 | 1400 | 1400 | 1460 | 1300 | 1822.951 | 1310.5295 | 1420 | 1335
| Std Dev. |  | 40.8248 | 51.6398 | 51.6398 | 51.6398 | 0 | 40.8248 | 51.6398 | 0 | 40.8248 | 81.6497 | 38.3255 | 51.6398 | 40.8248 | 63.2456 | 40.8248 | 22.5093 | 40.8248 | 47.195 | 8.7911 | 19.7222 | 12.1106
| Rel.Std.Dev. |  | 2.951% | 3.603% | 4.426% | 4.187% | 0% | 3.181% | 4.556% | 0% | 2.692% | 6.124% | 3.285% | 3.521% | 3.181% | 4.518% | 2.951% | 1.535% | 3.101% | 2.593% | 0.671% | 1.38% | 0.9083%
| PDM3 |  | 4.676% | 8.459% | -11.72% | -6.675% | -9.197% | -2.891% | -14.24% | -9.197% | 14.77% | 0.8924% | -11.71% | 10.98% | -2.891% | 5.937% | 4.676% | 10.98% | -0.3688% | 37.72% | -0.8574% | 8.144% | 0.8924%
| Z-Score (Absolute) |  | 0.34 | 0.67 | 1.12 | 0.67 | 0.9 | 0.34 | 1.35 | 0.9 | 1.24 | 0 | 1.12 | 0.9 | 0.34 | 0.45 | 0.34 | 0.9 | 0.11 | 3.28 | 0.16 | 0.65 | 0
| Analytical results for Mn in OREAS 605b (Certified Value 87 ppm) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 1
| 1 | 1 | 76.8741 | 90 | 86 | 126 | 82 | 87 | 90 | 88 | 82 | 87 | 91 | 88.6215 | 94 | 85 | 90 | 84 | 88.3 | 83 | 93.151 | 85.4246 | 88.67 | 87.591 |  | 1
| 1 | 2 | 77.8386 | 88 | 86 | 126 | 80 | 87 | 88 | 87 | 82 | 88 | 93 | 90.0939 | 96 | 82 | 90 | 84 | 89.1 | 79 | 90.923 | 86.4794 | 88.97 | 91.571 |  | 30
| 1 | 3 | 79.743 | 90 | 91 | 126 | 81 | 87 | 88 | 83 | 83 | 89 | 93 | 88.8659 | 95 | 81 | 90 | 85 | 87.8 | 81 | 88.979 | 87.4249 | 87.96 | 89.401 |  | 16
| 1 | 4 | 81.3138 | 88 | 92 | 125 | 83 | 87 | 87 | 83 | 84 | 88 | 98 | 89.5442 | 96 | 83 | 97 | 85 | 89.1 | 83 | 91.239 | 85.3854 | 88.81 | 87.021 |  | 86.9007
| 1 | 5 | 78.1554 | 86 | 95 | 125 | 82 | 86 | 89 | 79 | 83 | 88 | 92 | 88.4539 | 94 | 76 | 94 | 83 | 89.7 | 79 | 86.901 | 87.9525 | 87.23 | 88.178 |  | 92
| 1 | 6 | 82.2021 | 87 | 89 | 126 | 79 | 86 | 80 | 82 | 83 | 87 | 88 | 89.6038 | 96 | 83 | 91 | 84 | 92.2 | 82 | 86.365 | 86.8712 | 86.5 | 87.763
| Mean |  | 79.3545 | 88.1667 | 89.8333 | 125.6667 | 81.1667 | 86.6667 | 87 | 83.6667 | 82.8333 | 87.8333 | 92.5 | 89.1972 | 95.1667 | 81.6667 | 92 | 84.1667 | 89.3667 | 81.1667 | 89.593 | 86.5897 | 88.0233 | 88.5875
| Median |  | 78.9492 | 88 | 90 | 126 | 81.5 | 87 | 88 | 83 | 83 | 88 | 92.5 | 89.2051 | 95.5 | 82.5 | 90.5 | 84 | 89.1 | 81.5 | 89.951 | 86.6753 | 88.315 | 87.9705
| Std Dev. |  | 2.097 | 1.6021 | 3.5449 | 0.5164 | 1.472 | 0.5164 | 3.5777 | 3.3267 | 0.75277 | 0.75277 | 3.2711 | 0.64552 | 0.98319 | 3.0768 | 2.8983 | 0.75277 | 1.541 | 1.8348 | 2.6526 | 1.0444 | 0.98865 | 1.6648
| Rel.Std.Dev. |  | 2.643% | 1.817% | 3.946% | 0.4109% | 1.814% | 0.5958% | 4.112% | 3.976% | 0.9088% | 0.857% | 3.536% | 0.7237% | 1.033% | 3.768% | 3.15% | 0.8944% | 1.724% | 2.261% | 2.961% | 1.206% | 1.123% | 1.879%
| PDM3 |  | -8.684% | 1.457% | 3.375% | 44.61% | -6.598% | -0.2693% | 0.1142% | -3.722% | -4.68% | 1.073% | 6.443% | 2.643% | 9.512% | -6.023% | 5.868% | -3.146% | 2.838% | -6.598% | 3.098% | -0.3579% | 1.292% | 1.941%
| Z-Score (Absolute) |  | 2.04 | 0.06 | 0.45 | 8.98 | 1.61 | 0.3 | 0.22 | 1.01 | 1.21 | 0.02 | 1.09 | 0.3 | 1.72 | 1.49 | 0.97 | 0.9 | 0.34 | 1.61 | 0.4 | 0.32 | 0.02 | 0.16
| Analytical results for Mo in OREAS 605b (Certified Value 16.7 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 2
| 1 | 1 | 15.9 | 15.92 | 14 | 17.05 | 16.3 | 19 | 17.15 | 17.35 | 16.5 | 17.9 | 16.3164 | 17.95 | 16.3 | 18 | 16.12 | 16.9 | 15 | 17.638 | 16.9729 | 16.36 | 15.418 |  |  | 1
| 1 | 2 | 16.4 | 16.04 | 14 | 17.45 | 17.3 | 18.25 | 17 | 17.25 | 17 | 18.2 | 16.7368 | 19.49 | 15.9 | 19 | 16.47 | 16.3 | 15 | 17.242 | 13.9241 | 16.62 | 15.421 |  |  | 31
| 1 | 3 | 15.8 | 16.86 | 14 | 17.45 | 16.6 | 19.2 | 19.05 | 16.95 | 16.9 | 18.2 | 17.473 | 20.82 | 16.1 | 19 | 16.52 | 16.7 | 15 | 17.094 | 15.1827 | 16.38 | 15.019 |  |  | 16
| 1 | 4 | 16.2 | 17.33 | 14 | 16.9 | 15.25 | 18.45 | 17 | 17.6 | 16.8 | 19.7 | 17.3626 | 18.81 | 16.5 | 20 | 16.37 | 16.6 | 15 | 17.6242 | 15.5359 | 16.81 | 15.491 |  |  | 16.7174
| 1 | 5 | 15.8 | 17.64 | 14 | 17.15 | 14.4 | 17.4 | 18.7 | 16.35 | 16.8 | 18.2 | 17.0185 | 19.94 | 15.8 | 19 | 16.2 | 16.4 | 15 | 17.4762 | 14.5194 | 16.39 | 15.397 |  |  | 93
| 1 | 6 | 16.4 | 17.43 | 14 | 16.6 | 15.75 | 17.2 | 18.65 | 16.55 | 16.8 | 17.2 | 16.7956 | 18.02 | 16.1 | 19 | 16.45 | 17 | 14 | 16.9616 | 14.1252 | 16 | 15.682
| Mean |  | 16.0833 | 16.87 | 14 | 17.1 | 15.9333 | 18.25 | 17.925 | 17.0083 | 16.8 | 18.2333 | 16.9505 | 19.1717 | 16.1167 | 19 | 16.355 | 16.65 | 14.8333 | 17.3393 | 15.0434 | 16.4267 | 15.4047
| Median |  | 16.05 | 17.095 | 14 | 17.1 | 16.025 | 18.35 | 17.9 | 17.1 | 16.8 | 18.2 | 16.907 | 19.15 | 16.1 | 19 | 16.41 | 16.65 | 15 | 17.3591 | 14.8511 | 16.385 | 15.4195
| Std Dev. |  | 0.28577 | 0.73615 | 0 | 0.32863 | 1.0294 | 0.81609 | 0.96992 | 0.48417 | 0.16733 | 0.8165 | 0.42876 | 1.1277 | 0.25626 | 0.63246 | 0.16059 | 0.27386 | 0.40825 | 0.28333 | 1.1274 | 0.27362 | 0.21615
| Rel.Std.Dev. |  | 1.777% | 4.364% | 0% | 1.922% | 6.461% | 4.472% | 5.411% | 2.847% | 0.996% | 4.478% | 2.529% | 5.882% | 1.59% | 3.329% | 0.9819% | 1.645% | 2.752% | 1.634% | 7.494% | 1.666% | 1.403%
| PDM3 |  | -3.793% | 0.9127% | -16.26% | 2.288% | -4.69% | 9.168% | 7.223% | 1.74% | 0.494% | 9.068% | 1.394% | 14.68% | -3.594% | 13.65% | -2.168% | -0.4033% | -11.27% | 3.72% | -10.01% | -1.739% | -7.853%
| Z-Score (Absolute) |  | 0.83 | 0.04 | NA | 0.29 | 0.99 | 1.56 | 1.2 | 0.19 | 0.04 | 1.54 | 0.13 | 2.58 | 0.79 | NA | 0.53 | 0.2 | NA | 0.56 | 1.98 | 0.45 | 1.58
Results from laboratories 4, 19 & 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 605b (Certified Value 0.051 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  |  | 3
| 1 | 1 | 0.05 | 0.045 | 0.05 | 0.05 | 0.05 | 0.07 | 0.05 | 0.04 | 0.05 | 0.04 | 0.043953 | 0.051 | 0.04 | 0.05 | 0.0648 | 0.05 | 0.054723 | 0.059684 | 0.0619 | 0.048 |  |  |  | 1
| 1 | 2 | 0.05 | 0.046 | 0.04 | 0.05 | 0.05 | 0.06 | 0.04 | 0.04 | 0.05 | 0.05 | 0.044944 | 0.052 | 0.04 | 0.04 | 0.0637 | 0.05 | 0.055326 | 0.063184 | 0.06232 | 0.047 |  |  |  | #N/A
| 1 | 3 | 0.05 | 0.049 | 0.04 | 0.05 | 0.05 | 0.06 | 0.06 | 0.04 | 0.05 | 0.04 | 0.044065 | 0.052 | 0.04 | 0.05 | 0.062 | 0.05 | 0.054446 | 0.062083 | 0.06175 | 0.048 |  |  |  | 16
| 1 | 4 | 0.05 | 0.049 | 0.04 | 0.05 | 0.05 | 0.06 | 0.06 | 0.04 | 0.05 | 0.05 | 0.044675 | 0.053 | 0.04 | 0.04 | 0.0673 | 0.05 | 0.055845 | 0.06172 | 0.06222 | 0.048 |  |  |  | 0.050692
| 1 | 5 | 0.05 | 0.049 | 0.04 | 0.05 | 0.05 | 0.05 | 0.06 | 0.04 | 0.05 | 0.05 | 0.043543 | 0.052 | 0.04 | 0.05 | 0.0667 | 0.05 | 0.054978 | 0.063957 | 0.0618 | 0.048 |  |  |  | 94
| 1 | 6 | 0.05 | 0.048 | 0.04 | 0.05 | 0.05 | 0.05 | 0.06 | 0.04 | 0.05 | 0.05 | 0.044748 | 0.051 | 0.04 | 0.05 | 0.0646 | 0.05 | 0.054504 | 0.061294 | 0.06231 | 0.048
| Mean |  | 0.05 | 0.047667 | 0.041667 | 0.05 | 0.05 | 0.058333 | 0.055 | 0.04 | 0.05 | 0.046667 | 0.044321 | 0.051833 | 0.04 | 0.046667 | 0.06485 | 0.05 | 0.05497 | 0.061987 | 0.06205 | 0.047833
| Median |  | 0.05 | 0.0485 | 0.04 | 0.05 | 0.05 | 0.06 | 0.06 | 0.04 | 0.05 | 0.05 | 0.04437 | 0.052 | 0.04 | 0.05 | 0.0647 | 0.05 | 0.054851 | 0.061901 | 0.06206 | 0.048
| Std Dev. |  | 7.6012e-18 | 0.0017512 | 0.0040825 | 7.6012e-18 | 7.6012e-18 | 0.0075277 | 0.0083666 | 0 | 7.6012e-18 | 0.005164 | 0.00054815 | 0.00075277 | 0 | 0.005164 | 0.001946 | 7.6012e-18 | 0.00053762 | 0.0014952 | 0.00026245 | 0.00040825
| Rel.Std.Dev. |  | 1.52e-14% | 3.674% | 9.798% | 1.52e-14% | 1.52e-14% | 12.9% | 15.21% | 0% | 1.52e-14% | 11.07% | 1.237% | 1.452% | 0% | 11.07% | 3.001% | 1.52e-14% | 0.978% | 2.412% | 0.423% | 0.8535%
| PDM3 |  | -1.366% | -5.969% | -17.8% | -1.366% | -1.366% | 15.07% | 8.498% | -21.09% | -1.366% | -7.941% | -12.57% | 2.251% | -21.09% | -7.941% | 27.93% | -1.366% | 8.439% | 22.28% | 22.41% | -5.64%
| Z-Score (Absolute) |  | 0 | 0.38 | 1.35 | 0 | 0 | 1.35 | 0.81 | 1.62 | 0 | 0.54 | 0.92 | 0.3 | 1.62 | 0.54 | 2.41 | 0 | 0.81 | 1.95 | 1.96 | 0.35
| Analytical results for Nb in OREAS 605b (Certified Value 0.75 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | [6 empty] | 2
| 1 | 1 | 0.7 | 0.39 | < 5 | 0.69 | 0.72 | 0.68 | 0.57 | 0.62 | 0.99 | 1.35 | 0.87144 | 0.75 | 1.1 | 0.65 | 0.807 | 0.377 | < 1 | [6 empty] | 1
| 1 | 2 | 0.8 | 0.32 | < 5 | 0.68 | 0.7 | 0.59 | 0.56 | 0.64 | 1.01 | 1.4 | 1.015 | 0.78 | 1.09 | 0.65 | 0.733 | 0.385 | < 1 | [6 empty] | 32
| 1 | 3 | 0.6 | 0.31 | < 5 | 0.71 | 0.74 | 0.63 | 0.66 | 0.63 | 1.03 | 1.34 | 0.95599 | 0.96 | 1.08 | 0.67 | 0.822 | 0.375 | < 1 | [6 empty] | 16
| 1 | 4 | 0.7 | 0.36 | < 5 | 0.69 | 0.81 | 0.6 | 0.62 | 0.67 | 1.06 | 1.45 | 0.98625 | 0.74 | 1.12 | 0.66 | 0.729 | 0.37 | < 1 | [6 empty] | 0.75314
| 1 | 5 | 0.6 | 0.3 | < 5 | 0.7 | 0.78 | 0.64 | 0.64 | 0.61 | 0.99 | 1.31 | 0.93405 | 0.89 | 1.03 | 0.64 | 0.822 | 0.419 | < 1 | [6 empty] | 95
| 1 | 6 | 0.7 | 0.38 | < 5 | 0.67 | 0.83 | 0.63 | 0.63 | 0.6 | 0.97 | 1.35 | 0.92285 | 0.72 | 1.09 | 0.66 | 0.83 | 0.371 | < 1
| Mean |  | 0.68333 | 0.34333 |  | 0.69 | 0.76333 | 0.62833 | 0.61333 | 0.62833 | 1.0083 | 1.3667 | 0.94759 | 0.80667 | 1.085 | 0.655 | 0.7905 | 0.38283
| Median |  | 0.7 | 0.34 |  | 0.69 | 0.76 | 0.63 | 0.625 | 0.625 | 1 | 1.35 | 0.94502 | 0.765 | 1.09 | 0.655 | 0.8145 | 0.376
| Std Dev. |  | 0.075277 | 0.038297 |  | 0.014142 | 0.05164 | 0.031885 | 0.039833 | 0.024833 | 0.032506 | 0.050067 | 0.050401 | 0.096264 | 0.030166 | 0.010488 | 0.046702 | 0.018509
| Rel.Std.Dev. |  | 11.02% | 11.15% |  | 2.05% | 6.765% | 5.075% | 6.495% | 3.952% | 3.224% | 3.663% | 5.319% | 11.93% | 2.78% | 1.601% | 5.908% | 4.835%
| PDM3 |  | -9.269% | -54.41% |  | -8.384% | 1.353% | -16.57% | -18.56% | -16.57% | 33.88% | 81.46% | 25.82% | 7.107% | 44.06% | -13.03% | 4.96% | -49.17%
| Z-Score (Absolute) |  | NA | 1.55 | 7.15 | 0.15 | 0.15 | 0.4 | 0.46 | 0.4 | 1.14 | 2.58 | 0.89 | 0.32 | 1.45 | 0.29 | 0.26 | 1.39 | 0.91
Results from laboratory 2 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 605b (Indicative Value 5.78 ppm) | [24 empty] | Indicative
Nd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab | [6 empty] | 1
| No | No | 02: 03; 03: 10; 04: 23 | [6 empty] | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS | [6 empty] | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g | [6 empty] | 2
| 1 | 1 | 02: 5.69; 03: 6.36; 04: 5.706 | [6 empty] | 1
| 1 | 2 | 02: 5.44; 03: 6.59; 04: 5.792 | [6 empty] | 13
| 1 | 3 | 02: 6.31; 03: 6.42; 04: 5.694 | [6 empty] | 16
| 1 | 4 | 02: 5.78; 03: 6.44; 04: 5.873 | [6 empty] | 5.7766
| 1 | 5 | 02: 5.96; 03: 6.65; 04: 5.844 | [6 empty] | 19
| 1 | 6 | 02: 5.61; 03: 6.48; 04: 5.62
| Mean |  | 02: 5.7983; 03: 6.49; 04: 5.7548
| Median |  | 02: 5.735; 03: 6.46; 04: 5.749
| Std Dev. |  | 02: 0.30459; 03: 0.10954; 04: 0.09753
| Rel.Std.Dev. |  | 02: 5.253%; 03: 1.688%; 04: 1.695%
| PDM3 |  | 02: 0.3765%; 03: 12.35%; 04: -0.3765%
| Z-Score (Absolute) |  | 02: 0; 03: 10.72; 04: 0.67
| Analytical results for Ni in OREAS 605b (Certified Value 20.8 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 2
| 1 | 1 | 19.926 | 21 | 20.8 | 20 | 20.8 | 20.2 | 22.4 | 22.3 | 21.7 | 19.1 | 23 | 22.0812 | 18.5 | 19.6 | 28 | 19.9 | 22 | 19 | 21.835 | 21.2127 | 19.78 | 22.296 |  | 1
| 1 | 2 | 19.3557 | 21 | 20.9 | 20 | 20.7 | 20.5 | 21.2 | 21.8 | 21 | 20 | 23.4 | 22.4353 | 17.9 | 19.9 | 28 | 19.8 | 21.8 | 19 | 22.288 | 19.7553 | 20.07 | 22.395 |  | 34
| 1 | 3 | 19.3056 | 21 | 22.2 | 21 | 21.3 | 19.7 | 22 | 24.7 | 21.4 | 19.6 | 23.1 | 21.9225 | 19.2 | 19.8 | 27 | 20.3 | 21.8 | 19 | 21.183 | 19.9537 | 21.79 | 22.516 |  | 16
| 1 | 4 | 19.584 | 20 | 22.5 | 20 | 20.4 | 19.7 | 20.9 | 20.5 | 21.8 | 19.2 | 24.5 | 22.8661 | 17.1 | 19.8 | 28 | 20.1 | 21.7 | 20 | 21.961 | 20.1617 | 21.47 | 22.529 |  | 20.8224
| 1 | 5 | 19.9711 | 20 | 23 | 19 | 21.5 | 19.2 | 21.3 | 22.5 | 20.6 | 19.3 | 23.5 | 22.3514 | 19.4 | 20.1 | 30 | 19.7 | 21.8 | 20 | 22.031 | 20.3781 | 20.86 | 22.783 |  | 96
| 1 | 6 | 19.1404 | 21 | 22.6 | 20 | 20.2 | 21.3 | 20.5 | 22.1 | 20.6 | 19.6 | 22 | 22.6517 | 18 | 18.6 | 34 | 19.7 | 22.6 | 18 | 21.191 | 20.3175 | 19.78 | 21.84
| Mean |  | 19.5471 | 20.6667 | 22 | 20 | 20.8167 | 20.1 | 21.3833 | 22.3167 | 21.1833 | 19.4667 | 23.25 | 22.3847 | 18.35 | 19.6333 | 29.1667 | 19.9167 | 21.95 | 19.1667 | 21.7482 | 20.2965 | 20.625 | 22.3932
| Median |  | 19.4699 | 21 | 22.35 | 20 | 20.75 | 19.95 | 21.25 | 22.2 | 21.2 | 19.45 | 23.25 | 22.3933 | 18.25 | 19.8 | 28 | 19.85 | 21.8 | 19 | 21.898 | 20.2396 | 20.465 | 22.4555
| Std Dev. |  | 0.34209 | 0.5164 | 0.92736 | 0.63246 | 0.50365 | 0.74027 | 0.70261 | 1.3659 | 0.53072 | 0.33267 | 0.81179 | 0.34999 | 0.86429 | 0.53166 | 2.5626 | 0.24014 | 0.33317 | 0.75277 | 0.45915 | 0.50498 | 0.87893 | 0.31653
| Rel.Std.Dev. |  | 1.75% | 2.499% | 4.215% | 3.162% | 2.419% | 3.683% | 3.286% | 6.121% | 2.505% | 1.709% | 3.492% | 1.564% | 4.71% | 2.708% | 8.786% | 1.206% | 1.518% | 3.928% | 2.111% | 2.488% | 4.261% | 1.413%
| PDM3 |  | -6.125% | -0.7481% | 5.655% | -3.95% | -0.02774% | -3.47% | 2.694% | 7.176% | 1.733% | -6.511% | 11.66% | 7.503% | -11.87% | -5.711% | 40.07% | -4.35% | 5.415% | -7.952% | 4.446% | -2.526% | -0.9482% | 7.543%
| Z-Score (Absolute) |  | 0.7 | 0.04 | 0.74 | 0.43 | 0.04 | 0.38 | 0.38 | 0.92 | 0.26 | 0.75 | 1.47 | 0.96 | 1.4 | 0.65 | 4.93 | 0.48 | 0.71 | 0.92 | 0.59 | 0.26 | 0.07 | 0.97
| Analytical results for P in OREAS 605b (Certified Value 187 ppm) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 0
| 1 | 1 | 183 | 170 | 600 | 180 | 190 | 190 | 200 | 180 | 240 | 171 | 183.0989 | 190 | 180 | 178 | 186 | 200 | 300 | 222.163 | 180.2818 | < 10 | 680 |  |  | 1
| 1 | 2 | 178 | 160 | 600 | 180 | 180 | 190 | 200 | 180 | 240 | 161 | 216.698 | 200 | 180 | 177 | 191 | 197 | 300 | 239.891 | 180.1595 | < 10 | 780 |  |  | #N/A
| 1 | 3 | 179 | 170 | 600 | 190 | 190 | 200 | 180 | 180 | 250 | 173 | 185.8556 | 190 | 180 | 174 | 198 | 200 | 300 | 233.505 | 186.1272 | < 10 | 680 |  |  | 16
| 1 | 4 | 180 | 180 | 600 | 190 | 200 | 190 | 170 | 190 | 250 | 159 | 214.7672 | 200 | 180 | 186 | 186 | 197 | 200 | 229.105 | 179.7199 | < 10 | 680 |  |  | 187.1444
| 1 | 5 | 181 | 180 | 600 | 180 | 190 | 190 | 160 | 180 | 250 | 178 | 187.7948 | 190 | 180 | 188 | 192 | 200 | 300 | 211.875 | 185.6368 | < 10 | 730 |  |  | 97
| 1 | 6 | 182 | 180 | 600 | 180 | 200 | 180 | 160 | 180 | 230 | 152 | 211.7647 | 190 | 180 | 174 | 185 | 206 | 200 | 226.001 | 184.7746 | < 10 | 790
| Mean |  | 180.5 | 173.3333 | 600 | 183.3333 | 191.6667 | 190 | 178.3333 | 181.6667 | 243.3333 | 165.6667 | 199.9965 | 193.3333 | 180 | 179.5 | 189.6667 | 200 | 266.6667 | 227.09 | 182.7833 |  | 723.3333
| Median |  | 180.5 | 175 | 600 | 180 | 190 | 190 | 175 | 180 | 245 | 166 | 199.7798 | 190 | 180 | 177.5 | 188.5 | 200 | 300 | 227.553 | 182.5282 |  | 705
| Std Dev. |  | 1.8708 | 8.165 | 0 | 5.164 | 7.5277 | 6.3246 | 18.3485 | 4.0825 | 8.165 | 9.8725 | 15.9373 | 5.164 | 0 | 6.0581 | 5.0067 | 3.2863 | 51.6398 | 9.6544 | 3.0271 |  | 51.6398
| Rel.Std.Dev. |  | 1.036% | 4.711% | 0% | 2.817% | 3.928% | 3.329% | 10.29% | 2.247% | 3.355% | 5.959% | 7.969% | 2.671% | 0% | 3.375% | 2.64% | 1.643% | 19.36% | 4.251% | 1.656% |  | 7.139%
| PDM3 |  | -3.55% | -7.38% | 220.6% | -2.036% | 2.416% | 1.526% | -4.708% | -2.927% | 30.02% | -11.48% | 6.868% | 3.307% | -3.818% | -4.085% | 1.348% | 6.869% | 42.49% | 21.34% | -2.33% |  | 286.5%
| Z-Score (Absolute) |  | 0.53 | 1.17 | 36.77 | 0.28 | 0.46 | 0.31 | 0.73 | 0.43 | 5.05 | 1.85 | 1.2 | 0.61 | 0.58 | 0.62 | 0.28 | 1.2 | NA | 3.61 | 0.33 | 16.14 | 47.74
Results from laboratory 22 were removed due to their 100 ppm reading resolution.
| Analytical results for Pb in OREAS 605b (Certified Value 894 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 0
| 1 | 1 | 963.3 | 871.79 | 860 | 894 | 882 | 920 | 946 | 872 | 868 | 930.2 | 933.0865 | 824.63 | 988 | 904 | 926.7 | 888 | 827 | 925.028 | 879.6249 | 1035 | 846.935 |  |  | 1
| 1 | 2 | 958.5 | 870.1 | 850 | 889 | 886 | 897 | 923 | 870 | 870 | 1001 | 938.7903 | 823.77 | 981 | 909 | 914.6 | 876 | 811 | 915.357 | 882.915 | 1024 | 858.917 |  |  | 35
| 1 | 3 | 958.1 | 866.86 | 850 | 908 | 897 | 911 | 881 | 869 | 867 | 977.1 | 931.3044 | 847.85 | 988 | 905 | 930.6 | 885 | 820 | 903.268 | 893.991 | 1024 | 845.221 |  |  | 16
| 1 | 4 | 937.9 | 874.82 | 860 | 908 | 895 | 909 | 883 | 890 | 868 | 911.6 | 945.5336 | 833.37 | 984 | 959 | 940.3 | 819 | 835 | 915.208 | 877.92 | 1043 | 864.802 |  |  | 894.4079
| 1 | 5 | 934.7 | 864.94 | 850 | 901 | 905 | 895 | 841 | 878 | 865 | 979.2 | 930.3189 | 821.81 | 971 | 949 | 924.8 | 804 | 836 | 898.676 | 898.3114 | 1025 | 859.421 |  |  | 98
| 1 | 6 | 932.3 | 855.57 | 860 | 870 | 904 | 867 | 860 | 876 | 867 | 934.6 | 948.7757 | 844.16 | 1010 | 913 | 913.4 | 900 | 790 | 889.966 | 887.8098 | 1012 | 875.196
| Mean |  | 947.4667 | 867.3467 | 855 | 895 | 894.8333 | 899.8333 | 889 | 875.8333 | 867.5 | 955.6167 | 937.9682 | 832.5983 | 987 | 923.1667 | 925.0667 | 862 | 819.8333 | 907.9172 | 886.762 | 1027.1667 | 858.4153
| Median |  | 948 | 868.48 | 855 | 897.5 | 896 | 903 | 882 | 874 | 867.5 | 955.85 | 935.9384 | 829 | 986 | 911 | 925.75 | 880.5 | 823.5 | 909.238 | 885.3624 | 1024.5 | 859.169
| Std Dev. |  | 13.9286 | 6.7512 | 5.4772 | 14.3944 | 9.3256 | 18.5733 | 39.1152 | 7.7567 | 1.6432 | 34.9357 | 7.7648 | 11.1777 | 12.8996 | 24.3016 | 10.1101 | 40.1447 | 17.3829 | 12.8883 | 8.1322 | 10.6474 | 11.2208
| Rel.Std.Dev. |  | 1.47% | 0.7784% | 0.6406% | 1.608% | 1.042% | 2.064% | 4.4% | 0.8856% | 0.1894% | 3.656% | 0.8278% | 1.343% | 1.307% | 2.632% | 1.093% | 4.657% | 2.12% | 1.42% | 0.9171% | 1.037% | 1.307%
| PDM3 |  | 5.932% | -3.026% | -4.406% | 0.0662% | 0.04757% | 0.6066% | -0.6046% | -2.077% | -3.008% | 6.843% | 4.87% | -6.911% | 10.35% | 3.215% | 3.428% | -3.623% | -8.338% | 1.51% | -0.8549% | 14.84% | -4.024%
| Z-Score (Absolute) |  | 1.17 | 0.61 | 0.89 | 0 | 0 | 0.11 | 0.13 | 0.42 | 0.61 | 1.36 | 0.96 | 1.39 | 2.06 | 0.63 | 0.67 | 0.73 | 1.67 | 0.29 | 0.18 | 2.95 | 0.81
| Analytical results for Pd in OREAS 605b (Indicative Value < 10 ppb) | [24 empty] | Indicative
Pd | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab |  |  | 1
| No | No | 02: 03; 03: 11; 04: 12; 05: 23 |  |  | ppb
|  |  | 02: AR*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS |  |  | 1
|  |  | 02: 0.5g; 03: 10g; 04: 01g; 05: 0.25g |  |  | 1
| 1 | 1 | 02: < 10; 03: 12; 04: < 10; 05: < 50 |  |  | 1
| 1 | 2 | 02: < 10; 03: 13; 04: < 10; 05: < 50 |  |  | 1
| 1 | 3 | 02: < 10; 03: < 10; 04: < 10; 05: < 50 |  |  | 16
| 1 | 4 | 02: < 10; 03: < 10; 04: < 10; 05: < 50 |  |  | < 10
| 1 | 5 | 02: < 10; 03: < 10; 04: < 10; 05: < 50 |  |  | 20
| 1 | 6 | 02: 22; 03: < 10; 04: < 10; 05: < 50
| Mean |  | 02: 22; 03: 12.5
| Median |  | 02: 22; 03: 12.5
| Std Dev. |  | 03: 0.70711
| Rel.Std.Dev. |  | 03: 5.657%
PDM3
| Z-Score (Absolute) |  | 02: 0.08; 03: 0.08; 04: 1.27; 05: 8.25
| Analytical results for Pr in OREAS 605b (Indicative Value 1.61 ppm) | [24 empty] | Indicative
Pr | Round | Replicate | 02: Lab; 03: Lab; 04: Lab |  |  | 1
| No | No | 02: 03; 03: 10; 04: 23 |  |  | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS |  |  | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g |  |  | 2
| 1 | 1 | 02: 1.3; 03: 1.83; 04: 1.519 |  |  | 1
| 1 | 2 | 02: 1.38; 03: 1.84; 04: 1.515 |  |  | 15
| 1 | 3 | 02: 1.54; 03: 1.85; 04: 1.538 |  |  | 16
| 1 | 4 | 02: 1.59; 03: 1.86; 04: 1.551 |  |  | 1.6112
| 1 | 5 | 02: 1.49; 03: 1.87; 04: 1.524 |  |  | 21
| 1 | 6 | 02: 1.43; 03: 1.85; 04: 1.524
| Mean |  | 02: 1.455; 03: 1.85; 04: 1.5285
| Median |  | 02: 1.46; 03: 1.85; 04: 1.524
| Std Dev. |  | 02: 0.10672; 03: 0.014142; 04: 0.013487
| Rel.Std.Dev. |  | 02: 7.335%; 03: 0.7644%; 04: 0.8824%
| PDM3 |  | 02: -9.693%; 03: 14.82%; 04: -5.131%
| Z-Score (Absolute) |  | 02: 0.67; 03: 2.95; 04: 0
| Analytical results for Pt in OREAS 605b (Indicative Value < 5 ppb) | [24 empty] | Indicative
Pt | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab |  |  | 1
| No | No | 02: 03; 03: 11; 04: 12; 05: 23 |  |  | ppb
|  |  | 02: AR*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS |  |  | 2
|  |  | 02: 0.5g; 03: 10g; 04: 01g; 05: 0.25g |  |  | 2
| 1 | 1 | 02: < 2; 03: < 5; 04: < 5; 05: < 10 |  |  | 1
| 1 | 2 | 02: 2; 03: < 5; 04: < 5; 05: < 10 |  |  | 1
| 1 | 3 | 02: 2; 03: < 5; 04: < 5; 05: < 10 |  |  | 16
| 1 | 4 | 02: 8; 03: < 5; 04: < 5; 05: < 10 |  |  | < 5
| 1 | 5 | 02: 3; 03: < 5; 04: < 5; 05: < 10 |  |  | 22
| 1 | 6 | 02: 4; 03: < 5; 04: < 5; 05: < 10
| Mean |  | 02: 3.8
| Median |  | 02: 3
| Std Dev. |  | 02: 2.49
| Rel.Std.Dev. |  | 02: 65.53%
PDM3
| Z-Score (Absolute) |  | 02: 0.67; 03: 0.67; 04: 0.67; 05: 3.37
| Analytical results for Rb in OREAS 605b (Certified Value 17.5 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | [8 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [8 empty] | 2
| 1 | 1 | 16.38 | 17 | 16.9 | 17.1 | 19.3 | 18.9 | 18.2 | 16.4 | 19.64 | 18.4234 | 15.7 | 16.8 | 16.7 | 18.7 | 15.119 | [8 empty] | 1
| 1 | 2 | 16.52 | 16.6 | 17.1 | 17.7 | 18.9 | 18.4 | 17.9 | 16.5 | 19.68 | 18.7286 | 17.6 | 16.1 | 16.6 | 18.5 | 15.468 | [8 empty] | 37
| 1 | 3 | 16.63 | 17.5 | 17.2 | 18 | 19.1 | 18.8 | 17.2 | 16.3 | 19.91 | 18.2938 | 20.2 | 16.6 | 17 | 18.6 | 15.328 | [8 empty] | 16
| 1 | 4 | 15.88 | 17.9 | 16.8 | 16 | 18.4 | 17.4 | 18.1 | 16.1 | 19.58 | 18.9638 | 16.9 | 16.5 | 16.9 | 18.9 | 15.42 | [8 empty] | 17.4651
| 1 | 5 | 16.11 | 18.2 | 17.2 | 15.3 | 18.7 | 19.2 | 16.8 | 16.6 | 19.91 | 18.4051 | 19 | 16.1 | 16.5 | 18.6 | 15.46 | [8 empty] | 99
| 1 | 6 | 16.71 | 17.6 | 16.7 | 16.8 | 18.3 | 18.6 | 17.4 | 16.3 | 18.79 | 18.7127 | 14.5 | 16.3 | 16.7 | 18.7 | 14.735
| Mean |  | 16.3717 | 17.4667 | 16.9833 | 16.8167 | 18.7833 | 18.55 | 17.6 | 16.3667 | 19.585 | 18.5879 | 17.3167 | 16.4 | 16.7333 | 18.6667 | 15.255
| Median |  | 16.45 | 17.55 | 17 | 16.95 | 18.8 | 18.7 | 17.65 | 16.35 | 19.66 | 18.5681 | 17.25 | 16.4 | 16.7 | 18.65 | 15.374
| Std Dev. |  | 0.32046 | 0.58538 | 0.2137 | 1.0226 | 0.392 | 0.6253 | 0.55498 | 0.17512 | 0.41361 | 0.25442 | 2.0952 | 0.28284 | 0.18619 | 0.13663 | 0.28595
| Rel.Std.Dev. |  | 1.957% | 3.351% | 1.258% | 6.081% | 2.087% | 3.371% | 3.153% | 1.07% | 2.112% | 1.369% | 12.1% | 1.725% | 1.113% | 0.7319% | 1.874%
| PDM3 |  | -6.26% | 0.009197% | -2.758% | -3.713% | 7.548% | 6.212% | 0.7726% | -6.289% | 12.14% | 6.429% | -0.8497% | -6.098% | -4.19% | 6.88% | -12.65%
| Z-Score (Absolute) |  | 0.67 | 0.11 | 0.24 | 0.36 | 1.05 | 0.88 | 0.2 | 0.68 | 1.62 | 0.91 | 0 | 0.65 | 0.42 | 0.96 | 1.47
| Analytical results for Re in OREAS 605b (Certified Value 10.6 ppb) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 23 | [9 empty] | ppb
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [9 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | [9 empty] | 1
| 1 | 1 | < 50 | 11 | 11 | 12 | 10 | 11 | 11 | 10 | 11 | < 50 | 9 | 12 | 12 | 12 | [9 empty] | 1
| 1 | 2 | < 50 | 8 | 11 | 11 | 8 | 12 | 10 | 10 | 12 | < 50 | 8 | 12 | 11 | 11 | [9 empty] | 8
| 1 | 3 | < 50 | 11 | 10 | 9 | 8 | 11 | 10 | 10 | 12 | < 50 | 11 | 11 | 11 | 13 | [9 empty] | 16
| 1 | 4 | < 50 | 9 | 11 | 11 | 9 | 9 | 11 | 10 | < 1 | < 50 | 10 | 12 | 10 | 13 | [9 empty] | 10.5972
| 1 | 5 | < 50 | 9 | 11 | 16 | 7 | 10 | 11 | 10 | 12 | < 50 | 10 | 12 | 11 | 11 | [9 empty] | 100
| 1 | 6 | < 50 | 11 | 10 | 9 | 11 | 10 | 10 | 10 | 11 | < 50 | 11 | 11 | 11 | 14
| Mean |  |  | 9.8333 | 10.6667 | 11.3333 | 8.8333 | 10.5 | 10.5 | 10 | 11.6 |  | 9.8333 | 11.6667 | 11 | 12.3333
| Median |  |  | 10 | 11 | 11 | 8.5 | 10.5 | 10.5 | 10 | 12 |  | 10 | 12 | 11 | 12.5
| Std Dev. |  |  | 1.3292 | 0.5164 | 2.582 | 1.472 | 1.0488 | 0.54772 | 0 | 0.54772 |  | 1.169 | 0.5164 | 0.63246 | 1.2111
| Rel.Std.Dev. |  |  | 13.52% | 4.841% | 22.78% | 16.66% | 9.989% | 5.216% | 0% | 4.722% |  | 11.89% | 4.426% | 5.75% | 9.819%
| PDM3 |  |  | -7.208% | 0.6553% | 6.946% | -16.64% | -0.9174% | -0.9174% | -5.636% | 9.463% |  | -7.208% | 10.09% | 3.801% | 16.38%
| Z-Score (Absolute) |  | 12.96 | 0.67 | 0.07 | 0.67 | 1.57 | 0.07 | 0.07 | 0.52 | 0.75 | 12.96 | 0.67 | 0.97 | 0.37 | 1.57
| Analytical results for Rh in OREAS 605b (Indicative Value 34.8 ppb) | [24 empty] | Indicative
Rh | Round | Replicate | 02: Lab | [8 empty] | 1
| No | No | 02: 23 | [8 empty] | ppb
|  |  | 02: AR*MS | [8 empty] | 1
|  |  | 02: 0.25g | [8 empty] | 1
| 1 | 1 | 02: 40 | [8 empty] | 1
| 1 | 2 | 02: 33 | [8 empty] | 5
| 1 | 3 | 02: 31 | [8 empty] | 16
| 1 | 4 | 02: 39 | [8 empty] | 34.8333
| 1 | 5 | 02: 31 | [8 empty] | 16
| 1 | 6 | 02: 35
| Mean |  | 02: 34.8333
| Median |  | 02: 34
| Std Dev. |  | 02: 3.92
| Rel.Std.Dev. |  | 02: 11.25%
| PDM3 |  | 02: 1.11e-13%
Z-Score (Absolute) | bar: 02
| Analytical results for Ru in OREAS 605b (Indicative Value < 10 ppb) | [24 empty] | Indicative
Ru | Round | Replicate | 02: Lab | [8 empty] | 1
| No | No | 02: 23 | [8 empty] | ppb
|  |  | 02: AR*MS | [8 empty] | 1
|  |  | 02: 0.25g | [8 empty] | 1
| 1 | 1 | 02: < 10 | [8 empty] | 1
| 1 | 2 | 02: < 10 | [8 empty] | 6
| 1 | 3 | 02: < 10 | [8 empty] | 16
| 1 | 4 | 02: < 10 | [8 empty] | < 10
| 1 | 5 | 02: < 10 | [8 empty] | 17
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for S in OREAS 605b (Certified Value 9.21 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*AAS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.4g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 3
| 1 | 1 | > 5 | 8.55 | 9.24 | 9.51 | 9.46 | 9.41 | 9.84 | 9.53 | > 5 | > 5 | 7.438 | 8.95 | 9.45 | 8.09 | > 5 | > 5 | > 5 | 9.8252 | 8.7585 | 9.295 | > 5 |  |  | 1
| 1 | 2 | > 5 | 8.65 | 9.22 | 9.39 | 9.41 | NSS | 9.59 | 9.47 | > 5 | > 5 | 8.0651 | 8.96 | 9.64 | 8.1 | > 5 | > 5 | > 5 | 9.8343 | 8.7995 | 9.312 | > 5 |  |  | #N/A
| 1 | 3 | > 5 | 9.2 | 9.22 | 9.6 | 9.55 | NSS | 8.9 | 9.39 | > 5 | > 5 | 7.6739 | 8.81 | 9.56 | 8.07 | > 5 | > 5 | > 5 | 9.6465 | 8.9822 | 9.309 | > 5 |  |  | 16
| 1 | 4 | > 5 | 9.13 | 9.27 | 9.54 | 9.42 | 9.2 | 8.97 | 9.62 | > 5 | > 5 | 8.0426 | 9.09 | 9.89 | 8.59 | > 5 | > 5 | > 5 | 9.8987 | 8.8276 | 9.259 | > 5 |  |  | 9.2141
| 1 | 5 | > 5 | 9.08 | 9.2 | 9.51 | 9.4 | 9.21 | 8.55 | 9.56 | > 5 | > 5 | 7.6202 | 8.91 | 9.99 | 8.5 | > 5 | > 5 | > 5 | 9.6411 | 8.6962 | 9.268 | > 5 |  |  | 101
| 1 | 6 | > 5 | 9.1 | 9.26 | 9.32 | 9.45 | 9.37 | 8.81 | 9.46 | > 5 | > 5 | 7.9732 | 8.89 | 10 | 8.14 | > 5 | > 5 | > 5 | 9.5432 | 8.7696 | 9.252 | > 5
| Mean |  |  | 8.9517 | 9.235 | 9.4783 | 9.4483 | 9.2975 | 9.11 | 9.505 |  |  | 7.8022 | 8.935 | 9.755 | 8.2483 |  |  |  | 9.7315 | 8.8056 | 9.2825
| Median |  |  | 9.09 | 9.23 | 9.51 | 9.435 | 9.29 | 8.935 | 9.5 |  |  | 7.8235 | 8.93 | 9.765 | 8.12 |  |  |  | 9.7359 | 8.7845 | 9.2815
| Std Dev. |  |  | 0.27723 | 0.026646 | 0.10342 | 0.054924 | 0.10813 | 0.4961 | 0.08167 |  |  | 0.26015 | 0.092898 | 0.23569 | 0.23267 |  |  |  | 0.14012 | 0.097151 | 0.026159
| Rel.Std.Dev. |  |  | 3.097% | 0.2885% | 1.091% | 0.5813% | 1.163% | 5.446% | 0.8592% |  |  | 3.334% | 1.04% | 2.416% | 2.821% |  |  |  | 1.44% | 1.103% | 0.2818%
| PDM3 |  |  | -2.849% | 0.2264% | 2.867% | 2.542% | 0.9047% | -1.13% | 3.157% |  |  | -15.32% | -3.029% | 5.87% | -10.48% |  |  |  | 5.615% | -4.434% | 0.7419%
| Z-Score (Absolute) |  | NA | 0.75 | 0.06 | 0.54 | 0.46 | 0.09 | 0.36 | 0.6 | NA | NA | 3.55 | 0.79 | 1.21 | 2.46 | NA | NA | NA | 1.15 | 1.1 | 0.06 | NA
| Analytical results for Sb in OREAS 605b (Certified Value 323 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.25g | 0.2g | 01g | 0.25g |  |  | 0
| 1 | 1 | 255.85 | 316.7 | 280 | 352 | 335 | 360 | 370 | 340 | 325 | 347.78 | 311.8908 | 309.36 | 363 | 382 | 274.85 | 280 | 304 | 338.6923 | 310.3413 | 402.8 | 252.76 |  |  | 1
| 1 | 2 | 252.88 | 285.84 | 290 | 347 | 335 | 347 | 358 | 342 | 321 | 355.48 | 314.7968 | 298.84 | 360 | 377 | 270.66 | 274 | 287 | 347.2068 | 304.0267 | 396.2 | 254.688 |  |  | 39
| 1 | 3 | 243.73 | 297.16 | 290 | 354 | 333 | 357 | 345 | 339 | 323 | 355.96 | 312.9679 | 305.12 | 359 | 390 | 278.37 | 281 | 295 | 344.9517 | 316.8566 | 397.5 | 250.813 |  |  | 16
| 1 | 4 | 248.25 | 306.01 | 290 | 354 | 356 | 355 | 358 | 346 | 322 | 387.54 | 315.1475 | 303.06 | 374 | 412 | 278.05 | 281 | 300 | 345.3772 | 319.9784 | 406 | 263.584 |  |  | 323.3045
| 1 | 5 | 247.92 | 300.87 | 280 | 350 | 360 | 351 | 332 | 343 | 317 | 359.7 | 310.8154 | 310.55 | 356 | 408 | 268.59 | 287 | 296 | 352.3669 | 315.8906 | 399.6 | 259.519 |  |  | 102
| 1 | 6 | 251.69 | 303.67 | 290 | 348 | 359 | 340 | 347 | 339 | 319 | 343.33 | 314.5046 | 333.73 | 361 | 392 | 273.74 | 293 | 273 | 335.8161 | 308.9462 | 395.8 | 264.885
| Mean |  | 250.0533 | 301.7083 | 286.6667 | 350.8333 | 346.3333 | 351.6667 | 351.6667 | 341.5 | 321.1667 | 358.2983 | 313.3538 | 310.11 | 362.1667 | 393.5 | 274.0433 | 282.6667 | 292.5 | 344.0685 | 312.6733 | 399.65 | 257.7082
| Median |  | 249.97 | 302.27 | 290 | 351 | 345.5 | 353 | 352.5 | 341 | 321.5 | 355.72 | 313.7362 | 307.24 | 360.5 | 391 | 274.295 | 281 | 295.5 | 345.1644 | 313.116 | 398.55 | 257.1035
| Std Dev. |  | 4.2909 | 10.2009 | 5.164 | 2.9944 | 13.2313 | 7.3121 | 13.1859 | 2.7386 | 2.8577 | 15.5251 | 1.7525 | 12.3311 | 6.2423 | 13.9392 | 3.9167 | 6.532 | 11.1131 | 5.9701 | 5.9213 | 4.0397 | 5.8388
| Rel.Std.Dev. |  | 1.716% | 3.381% | 1.801% | 0.8535% | 3.82% | 2.079% | 3.75% | 0.8019% | 0.8898% | 4.333% | 0.5593% | 3.976% | 1.724% | 3.542% | 1.429% | 2.311% | 3.799% | 1.735% | 1.894% | 1.011% | 2.266%
| PDM3 |  | -22.66% | -6.68% | -11.33% | 8.515% | 7.123% | 8.773% | 8.773% | 5.628% | -0.6612% | 10.82% | -3.078% | -4.081% | 12.02% | 21.71% | -15.24% | -12.57% | -9.528% | 6.422% | -3.288% | 23.61% | -20.29%
| Z-Score (Absolute) |  | 1.57 | 0.43 | 0.76 | 0.66 | 0.56 | 0.67 | 0.67 | 0.45 | 0 | 0.82 | 0.17 | 0.24 | 0.91 | 1.6 | 1.04 | 0.85 | 0.63 | 0.51 | 0.19 | 1.74 | 1.4
| Analytical results for Sc in OREAS 605b (Certified Value 1.86 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 2
| 1 | 1 | 2 | 1.7 | 94 | 1.9 | 1.9 | 1.9 | 1.8 | 1.7 | 3 | 1.8 | 2.1433 | 2.8 | 2 | 2 | 2.5 | 1.8 | 1.5 | 1.545 | < 5 | 2.238 | 1.503 |  |  | 1
| 1 | 2 | 2 | 1.8 | 93 | 1.9 | 1.8 | 1.9 | 1.8 | 1.7 | 3 | 1.8 | 2.0733 | 2.8 | 1.9 | 2.1 | 2.3 | 1.68 | 1.5 | 1.553 | < 5 | 2.252 | 1.56 |  |  | 40
| 1 | 3 | 2 | 1.7 | 94 | 1.9 | 1.8 | 1.9 | 2 | 1.6 | 3 | 1.8 | 1.9806 | 2.8 | 2 | 2 | 2.3 | 1.66 | 1.5 | 1.538 | < 5 | 2.114 | 1.528 |  |  | 16
| 1 | 4 | 2 | 1.7 | 92 | 1.9 | 1.8 | 1.8 | 1.9 | 1.7 | 3 | 1.9 | 2.0929 | 2.9 | 1.9 | 2.2 | 2.4 | 1.72 | 1.5 | 1.57 | < 5 | 2.322 | 1.582 |  |  | 1.8553
| 1 | 5 | 2 | 1.8 | 93 | 2 | 1.7 | 1.8 | 2.1 | 1.6 | 3 | 2 | 1.9767 | 2.7 | 2 | 2.1 | 2.4 | 1.7 | 1.5 | 1.53 | < 5 | 2.142 | 1.522 |  |  | 103
| 1 | 6 | 2 | 1.7 | 93 | 1.8 | 1.7 | 1.7 | 2.1 | 1.6 | 3 | 1.9 | 2.0538 | 2.8 | 1.9 | 2.1 | 2.4 | 1.77 | 1.4 | 1.546 | < 5 | 2.132 | 1.583
| Mean |  | 2 | 1.7333 | 93.1667 | 1.9 | 1.7833 | 1.8333 | 1.95 | 1.65 | 3 | 1.8667 | 2.0534 | 2.8 | 1.95 | 2.0833 | 2.3833 | 1.7217 | 1.4833 | 1.547 |  | 2.2 | 1.5463
| Median |  | 2 | 1.7 | 93 | 1.9 | 1.8 | 1.85 | 1.95 | 1.65 | 3 | 1.85 | 2.0636 | 2.8 | 1.95 | 2.1 | 2.4 | 1.71 | 1.5 | 1.5455 |  | 2.19 | 1.544
| Std Dev. |  | 0 | 0.05164 | 0.75277 | 0.063246 | 0.075277 | 0.08165 | 0.13784 | 0.054772 | 0 | 0.08165 | 0.065131 | 0.063246 | 0.054772 | 0.075277 | 0.075277 | 0.053821 | 0.040825 | 0.013711 |  | 0.082965 | 0.033494
| Rel.Std.Dev. |  | 0% | 2.979% | 0.808% | 3.329% | 4.221% | 4.454% | 7.069% | 3.32% | 0% | 4.374% | 3.172% | 2.259% | 2.809% | 3.613% | 3.158% | 3.126% | 2.752% | 0.8863% |  | 3.771% | 2.166%
| PDM3 |  | 7.798% | -6.575% | 4922% | 2.408% | -3.88% | -1.185% | 5.103% | -11.07% | 61.7% | 0.6116% | 10.68% | 50.92% | 5.103% | 12.29% | 28.46% | -7.204% | -20.05% | -16.62% |  | 18.58% | -16.65%
| Z-Score (Absolute) |  | NA | 0.55 | NA | 0.06 | 0.36 | 0.18 | 0.24 | 0.85 | NA | 0.06 | 0.62 | 3.34 | 0.24 | 0.73 | 1.82 | 0.59 | 1.46 | 1.23 | 2.25 | 1.15 | 1.23
Results from laboratories 2, 4 & 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 605b (Certified Value 84 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | [6 empty] | 1
| 1 | 1 | 77 | 84.4 | < 5 | 93.4 | 76.2 | 85.9 | 93.4 | 90.1 | 69 | 88 | 77.4663 | 68.6 | 97.9 | 86 | 77.6 | 89.3615 | 85.9393 | [6 empty] | 1
| 1 | 2 | 75 | 81.8 | < 5 | 93.7 | 84.8 | 87.3 | 91.3 | 85.7 | 72 | 91 | 85.1541 | 64.3 | 95.8 | 83 | 80.4 | 89.2463 | 83.0679 | [6 empty] | 41
| 1 | 3 | 76 | 85.5 | < 5 | 94.3 | 83.6 | 87.5 | 92.1 | 87.4 | 72 | 90 | 78.5199 | 72.4 | 97.3 | 85 | 80.8 | 92.4216 | 84.2434 | [6 empty] | 16
| 1 | 4 | 74 | 86.8 | < 5 | 94.4 | 73.1 | 80.5 | 88 | 88.4 | 73 | 89 | 83.8242 | 65.3 | 99.4 | 87 | 79.9 | 92.3244 | 82.6541 | [6 empty] | 83.975
| 1 | 5 | 75 | 87.6 | < 5 | 94.4 | 69.3 | 86 | 95 | 84.9 | 73 | 90 | 80.6454 | 71 | 94.1 | 84 | 78.4 | 91.1435 | 83.0998 | [6 empty] | 104
| 1 | 6 | 77 | 85.1 | < 5 | 90.2 | 83.2 | 83.5 | 94.1 | 86.3 | 71 | 88 | 82.1074 | 68.3 | 97.1 | 85 | 80.1 | 87.7388 | 80.9065
| Mean |  | 75.6667 | 85.2 |  | 93.4 | 78.3667 | 85.1167 | 92.3167 | 87.1333 | 71.6667 | 89.3333 | 81.2862 | 68.3167 | 96.9333 | 85 | 79.5333 | 90.3727 | 83.3185
| Median |  | 75.5 | 85.3 |  | 94 | 79.7 | 85.95 | 92.75 | 86.85 | 72 | 89.5 | 81.3764 | 68.45 | 97.2 | 85 | 80 | 90.2525 | 83.0839
| Std Dev. |  | 1.2111 | 2.0308 |  | 1.6211 | 6.4308 | 2.6746 | 2.5007 | 1.9086 | 1.5055 | 1.2111 | 2.991 | 3.1365 | 1.8162 | 1.4142 | 1.2517 | 1.8884 | 1.6789
| Rel.Std.Dev. |  | 1.601% | 2.384% |  | 1.736% | 8.206% | 3.142% | 2.709% | 2.19% | 2.101% | 1.356% | 3.68% | 4.591% | 1.874% | 1.664% | 1.574% | 2.09% | 2.015%
| PDM3 |  | -9.894% | 1.459% |  | 11.22% | -6.679% | 1.359% | 9.933% | 3.761% | -14.66% | 6.381% | -3.202% | -18.65% | 15.43% | 1.221% | -5.289% | 7.619% | -0.7818%
| Z-Score (Absolute) |  | 1.15 | 0.02 | 10.18 | 1.04 | 0.82 | 0.01 | 0.9 | 0.26 | 1.64 | 0.53 | 0.46 | 2.06 | 1.47 | 0 | 0.67 | 0.66 | 0.21
| Analytical results for Sm in OREAS 605b (Indicative Value 1.12 ppm) | [24 empty] | Indicative
Sm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab | [6 empty] | 1
| No | No | 02: 03; 03: 10; 04: 23 | [6 empty] | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS | [6 empty] | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g | [6 empty] | 2
| 1 | 1 | 02: 0.99; 03: 1.27; 04: 1.082 | [6 empty] | 1
| 1 | 2 | 02: 1.14; 03: 1.25; 04: 1.075 | [6 empty] | 12
| 1 | 3 | 02: 1.09; 03: 1.25; 04: 0.989 | [6 empty] | 16
| 1 | 4 | 02: 1.01; 03: 1.27; 04: 1.035 | [6 empty] | 1.1243
| 1 | 5 | 02: 1.17; 03: 1.29; 04: 1.018 | [6 empty] | 18
| 1 | 6 | 02: 1.04; 03: 1.29; 04: 0.978
| Mean |  | 02: 1.0733; 03: 1.27; 04: 1.0295
| Median |  | 02: 1.065; 03: 1.27; 04: 1.0265
| Std Dev. |  | 02: 0.072296; 03: 0.017889; 04: 0.043085
| Rel.Std.Dev. |  | 02: 6.736%; 03: 1.409%; 04: 4.185%
| PDM3 |  | 02: -4.531%; 03: 12.96%; 04: -8.43%
| Z-Score (Absolute) |  | 02: 0; 03: 3.03; 04: 0.67
| Analytical results for Sn in OREAS 605b (Certified Value 5.87 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.25g | 0.2g | 01g | 0.25g |  |  | 3
| 1 | 1 | 5.9 | 6.1 | 6 | 5.7 | 5.9 | 6.1 | 6.1 | 5.3 | 5.46 | 5.99 | 5.6325 | 5.1 | 5.9 | < 10 | 6.5 | 5.86 | < 10 | 5.6668 | < 20 | < 5 | < 10 |  |  | 1
| 1 | 2 | 6.1 | 6.4 | 6 | 5.9 | 6.3 | 5.8 | 6 | 5.4 | 5.55 | 6.1 | 5.7889 | 5 | 5.9 | < 10 | 6.1 | 5.73 | < 10 | 5.8536 | < 20 | < 5 | < 10 |  |  | 26
| 1 | 3 | 6 | 6 | 7 | 5.9 | 5.7 | 6.4 | 5.9 | 5.3 | 5.39 | 6.04 | 5.5721 | 5 | 5.8 | < 10 | 6.4 | 5.88 | < 10 | 5.9098 | < 20 | < 5 | < 10 |  |  | 16
| 1 | 4 | 6.1 | 6.2 | 6 | 5.8 | 5.9 | 6 | 5.5 | 5.5 | 5.45 | 6.18 | 5.7761 | 4.5 | 6 | 10 | 6.2 | 5.89 | < 10 | 5.8558 | < 20 | < 5 | < 10 |  |  | 5.8749
| 1 | 5 | 6.2 | 6.2 | 6 | 5.8 | 5.6 | 6.1 | 6.1 | 5.3 | 5.3 | 6.12 | 5.5831 | 4.9 | 5.8 | < 10 | 6.1 | 5.79 | < 10 | 5.8471 | < 20 | < 5 | < 10 |  |  | 105
| 1 | 6 | 6.2 | 6.3 | 7 | 5.6 | 5.8 | 6 | 6 | 5.1 | 5.34 | 5.92 | 5.721 | 5.1 | 5.9 | < 10 | 6.1 | 5.95 | < 10 | 5.6236 | < 20 | < 5 | < 10
| Mean |  | 6.0833 | 6.2 | 6.3333 | 5.7833 | 5.8667 | 6.0667 | 5.9333 | 5.3167 | 5.415 | 6.0583 | 5.6789 | 4.9333 | 5.8833 | 10 | 6.2333 | 5.85 |  | 5.7928
| Median |  | 6.1 | 6.2 | 6 | 5.8 | 5.85 | 6.05 | 6 | 5.3 | 5.42 | 6.07 | 5.6768 | 5 | 5.9 | 10 | 6.15 | 5.87 |  | 5.8504
| Std Dev. |  | 0.1169 | 0.14142 | 0.5164 | 0.1169 | 0.24221 | 0.19664 | 0.22509 | 0.13292 | 0.090499 | 0.094322 | 0.095979 | 0.22509 | 0.075277 |  | 0.17512 | 0.07823 |  | 0.11732
| Rel.Std.Dev. |  | 1.922% | 2.281% | 8.154% | 2.021% | 4.129% | 3.241% | 3.794% | 2.5% | 1.671% | 1.557% | 1.69% | 4.563% | 1.28% |  | 2.809% | 1.337% |  | 2.025%
| PDM3 |  | 3.548% | 5.534% | 7.804% | -1.558% | -0.1399% | 3.264% | 0.9949% | -9.502% | -7.828% | 3.123% | -3.335% | -16.03% | 0.1438% | 70.22% | 6.101% | -0.4236% |  | -1.398%
| Z-Score (Absolute) |  | 0.7 | 1.04 | NA | 0.17 | 0.07 | 0.65 | 0.26 | 1.52 | 1.23 | 0.63 | 0.47 | 2.62 | 0.12 | 0.02 | 1.13 | 0.02 | 2.43 | 0.14 | 12.02 | 9.66 | 2.43
Results from laboratory 4 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 605b (Certified Value 34.2 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 1
| 1 | 1 | 68.2 | 31.5 | 35 | 39.1 | 32.2 | 40.6 | 34.1 | 32.1 | 55.5 | 26.56 | 53.0672 | 34.7 | 44.1 | 21.4 | 31.6 | 36.2 | 38 | 52.304 | 66.4982 | 77.21 | 36.52 |  |  | 1
| 1 | 2 | 67.9 | 31.2 | 34 | 38.8 | 33.4 | 35.6 | 33.8 | 33.6 | 55.5 | 25.47 | 53.1286 | 31.7 | 43.9 | 20.3 | 33.3 | 35.2 | 37.2 | 55.856 | 68.6403 | 77.94 | 37.39 |  |  | 27
| 1 | 3 | 67.3 | 33.1 | 34 | 40.6 | 35 | 36.2 | 38.3 | 34.2 | 55.6 | 25.17 | 53.5852 | 37 | 42.8 | 19.2 | 32.6 | 33.3 | 35.8 | 55.214 | 70.5848 | 78.15 | 36.57 |  |  | 16
| 1 | 4 | 64.5 | 32.9 | 34 | 38.6 | 34.7 | 36.6 | 35.2 | 37.7 | 56.8 | 27.79 | 54.9803 | 31 | 43.4 | 21.6 | 32.3 | 38.4 | 35.7 | 57.96 | 69.35 | 79.97 | 39.14 |  |  | 34.1703
| 1 | 5 | 63.8 | 31.5 | 34 | 40 | 33 | 34.9 | 37.9 | 30.3 | 56.4 | 22.79 | 54.501 | 35.3 | 41 | 20.8 | 32.8 | 40.1 | 35 | 54.292 | 68.5701 | 77.57 | 39.21 |  |  | 106
| 1 | 6 | 68.6 | 33.2 | 35 | 37.6 | 34.9 | 35.5 | 38.2 | 34.2 | 56.7 | 26.98 | 55.109 | 32.5 | 43.5 | 23.7 | 32.8 | 36.5 | 35.4 | 51.136 | 69.6112 | 78.29 | 37.64
| Mean |  | 66.7167 | 32.2333 | 34.3333 | 39.1167 | 33.8667 | 36.5667 | 36.25 | 33.6833 | 56.0833 | 25.7933 | 54.0619 | 33.7 | 43.1167 | 21.1667 | 32.5667 | 36.6167 | 36.1833 | 54.4603 | 68.8758 | 78.1883 | 37.745
| Median |  | 67.6 | 32.2 | 34 | 38.95 | 34.05 | 35.9 | 36.55 | 33.9 | 56 | 26.015 | 54.0431 | 33.6 | 43.45 | 21.1 | 32.7 | 36.35 | 35.75 | 54.753 | 68.9951 | 78.045 | 37.515
| Std Dev. |  | 2.0449 | 0.92448 | 0.5164 | 1.0629 | 1.1656 | 2.0617 | 2.1192 | 2.475 | 0.61779 | 1.7604 | 0.91872 | 2.3315 | 1.1303 | 1.5108 | 0.57504 | 2.3878 | 1.16 | 2.4695 | 1.3776 | 0.95784 | 1.1925
| Rel.Std.Dev. |  | 3.065% | 2.868% | 1.504% | 2.717% | 3.442% | 5.638% | 5.846% | 7.348% | 1.102% | 6.825% | 1.699% | 6.918% | 2.622% | 7.138% | 1.766% | 6.521% | 3.206% | 4.535% | 2% | 1.225% | 3.159%
| PDM3 |  | 95.25% | -5.669% | 0.477% | 14.48% | -0.8887% | 7.013% | 6.086% | -1.425% | 64.13% | -24.52% | 58.21% | -1.376% | 26.18% | -38.06% | -4.693% | 7.159% | 5.891% | 59.38% | 101.6% | 128.8% | 10.46%
| Z-Score (Absolute) |  | 5.08 | 0.73 | 0.38 | 0.43 | 0.46 | 0 | 0.05 | 0.49 | 3.29 | 1.82 | 2.95 | 0.48 | 1.1 | 2.6 | 0.67 | 0.01 | 0.06 | 3.02 | 5.45 | 7.02 | 0.2
| Analytical results for Ta in OREAS 605b (Indicative Value < 0.01 ppm) | [24 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | [6 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.02 | < 10 | [6 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.02 | < 10 | [6 empty] | 13
| 1 | 3 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.02 | < 10 | [6 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 5 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.02 | < 10 | [6 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 5 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.02 | < 10 | [6 empty] | 19
| 1 | 6 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.05 | < 0.02 | < 10
| Mean |  | 06: 0.01
| Median |  | 06: 0.01
| Std Dev. |  | 06: 0
Rel.Std.Dev. | bar: 06
PDM3
| Z-Score (Absolute) |  | 2.02 | 2.02 | 335.81 | 0.67 | 0.45 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 2.02 | 2.02 | 0.67 | 2.02 | 2.02 | 0 | 672.96
| Analytical results for Tb in OREAS 605b (Indicative Value 0.13 ppm) | [24 empty] | Indicative
Tb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [6 empty] | 1
| No | No | 02: 03; 03: 10; 04: 20; 05: 21; 06: 23 | [6 empty] | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*OES | [6 empty] | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g; 05: 0.25g; 06: 0.5g | [6 empty] | 2
| 1 | 1 | 02: 0.13; 03: 0.13; 04: 0.16; 05: 0.128; 06: 0.682 | [6 empty] | 1
| 1 | 2 | 02: 0.1; 03: 0.13; 04: 0.16; 05: 0.133; 06: 0.796 | [6 empty] | 14
| 1 | 3 | 02: 0.11; 03: 0.13; 04: 0.16; 05: 0.132; 06: 0.604 | [6 empty] | 16
| 1 | 4 | 02: 0.11; 03: 0.13; 04: 0.15; 05: 0.13; 06: 0.715 | [6 empty] | 0.13487
| 1 | 5 | 02: 0.14; 03: 0.14; 04: 0.16; 05: 0.137; 06: 0.903 | [6 empty] | 20
| 1 | 6 | 02: 0.13; 03: 0.13; 04: 0.16; 05: 0.127; 06: 0.577
| Mean |  | 02: 0.12; 03: 0.13167; 04: 0.15833; 05: 0.13117; 06: 0.71283
| Median |  | 02: 0.12; 03: 0.13; 04: 0.16; 05: 0.131; 06: 0.6985
| Std Dev. |  | 02: 0.015492; 03: 0.0040825; 04: 0.0040825; 05: 0.003656; 06: 0.12193
| Rel.Std.Dev. |  | 02: 12.91%; 03: 3.101%; 04: 2.578%; 05: 2.787%; 06: 17.1%
| PDM3 |  | 02: -11.03%; 03: -2.379%; 04: 17.39%; 05: -2.749%; 06: 428.5%
| Z-Score (Absolute) |  | 02: 0.67; 03: 0; 04: 1.54; 05: 0.03; 06: 33.59
| Analytical results for Te in OREAS 605b (Certified Value 33.4 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [6 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | [6 empty] | 2
| 1 | 1 | 31.14 | 33.1 | 37 | 34.1 | 30.9 | 28.2 | 35.5 | 34.7 | 30.1 | 35.8 | 33.3643 | 27.01 | 34 | 36.98 | 32 | 31.625 | 33.6724 | [6 empty] | 1
| 1 | 2 | 31.04 | 32.66 | 37 | 34.8 | 33.9 | 27 | 34.7 | 34.6 | 30.5 | 35.5 | 32.3287 | 24.6 | 33.2 | 36.13 | 31.4 | 31.5359 | 35.3183 | [6 empty] | 30
| 1 | 3 | 31.41 | 33.72 | 38 | 34.9 | 32.3 | 28.2 | 31.8 | 34.5 | 30.3 | 35.1 | 32.1401 | 25.75 | 33.5 | 36.88 | 32.2 | 33.6765 | 36.2685 | [6 empty] | 16
| 1 | 4 | 31.78 | 34.19 | 38 | 34.4 | 34.1 | 27.4 | 29.6 | 35.5 | 29.9 | 36.3 | 32.5776 | 24.4 | 34.2 | 36.94 | 32.2 | 32.9527 | 37.9836 | [6 empty] | 33.3772
| 1 | 5 | 31.72 | 34.57 | 37 | 34.9 | 31.1 | 28.7 | 32.7 | 33.2 | 30 | 35.4 | 32.0231 | 26.27 | 32.9 | 36.05 | 32.3 | 32.9725 | 37.046 | [6 empty] | 107
| 1 | 6 | 32.16 | 32.93 | 37 | 34.2 | 36.5 | 31 | 32.5 | 33.5 | 30.2 | 34.1 | 33.0348 | 25.61 | 33.5 | 35.82 | 32.8 | 32.5765 | 36.1987
| Mean |  | 31.5417 | 33.5283 | 37.3333 | 34.55 | 33.1333 | 28.4167 | 32.8 | 34.3333 | 30.1667 | 35.3667 | 32.5781 | 25.6067 | 33.55 | 36.4667 | 32.15 | 32.5565 | 36.0812
| Median |  | 31.565 | 33.41 | 37 | 34.6 | 33.1 | 28.2 | 32.6 | 34.55 | 30.15 | 35.45 | 32.4532 | 25.68 | 33.5 | 36.505 | 32.2 | 32.7646 | 36.2336
| Std Dev. |  | 0.42457 | 0.75566 | 0.5164 | 0.36194 | 2.1295 | 1.4063 | 2.109 | 0.84538 | 0.21602 | 0.74207 | 0.527 | 0.98988 | 0.48477 | 0.52221 | 0.45497 | 0.83596 | 1.4821
| Rel.Std.Dev. |  | 1.346% | 2.254% | 1.383% | 1.048% | 6.427% | 4.949% | 6.43% | 2.462% | 0.7161% | 2.098% | 1.618% | 3.866% | 1.445% | 1.432% | 1.415% | 2.568% | 4.108%
| PDM3 |  | -5.499% | 0.4527% | 11.85% | 3.514% | -0.7308% | -14.86% | -1.729% | 2.865% | -9.619% | 5.96% | -2.394% | -23.28% | 0.5176% | 9.256% | -3.677% | -2.459% | 8.101%
| Z-Score (Absolute) |  | 0.76 | 0.19 | 2 | 0.67 | 0 | 2.25 | 0.16 | 0.57 | 1.41 | 1.06 | 0.26 | 3.58 | 0.2 | 1.59 | 0.47 | 0.27 | 1.4
| Analytical results for Th in OREAS 605b (Indicative Value 2.01 ppm) | [24 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | 01g | [6 empty] | 2
| 1 | 1 | 3.56 | 1.2 | 1.5 | 1.3 | 1.5 | 1.2 | 1.1 | 2.72 | 1.08 | 2.8428 | 1.5 | 3.1 | 1.8 | 2.24 | 4.5979 | < 10 | 3.39 | [6 empty] | 1
| 1 | 2 | 3.61 | 1.3 | 1.4 | 1.4 | 1.2 | 1.2 | 1.1 | 2.67 | 1.12 | 2.9101 | 1.5 | 3 | 1.7 | 2.13 | 4.7082 | < 10 | 3.33 | [6 empty] | 15
| 1 | 3 | 3.74 | 1.3 | 1.5 | 1.3 | 1.3 | 1.3 | 1.2 | 2.69 | 1.03 | 2.9427 | 1.6 | 3.1 | 1.8 | 1.95 | 4.8143 | < 10 | 3.77 | [6 empty] | 16
| 1 | 4 | 3.56 | 1.3 | 1.4 | 1.5 | 1.4 | 1.1 | 1.3 | 2.67 | 1.03 | 3.0207 | 1.5 | 3 | 1.8 | 2.43 | 4.7492 | < 10 | 3.52 | [6 empty] | 2.0075
| 1 | 5 | 3.56 | 1.2 | 1.4 | 1.4 | 1.3 | 1.2 | 1.1 | 2.77 | 0.96 | 3.0572 | 1.5 | 3 | 1.8 | 2.56 | 4.7386 | < 10 | 3.42 | [6 empty] | 21
| 1 | 6 | 3.63 | 1.2 | 1.4 | 1.5 | 1.3 | 1.3 | 1.2 | 2.73 | 1.07 | 2.9705 | 1.6 | 3.1 | 1.9 | 2.15 | 4.6347 | < 10 | 3.28
| Mean |  | 3.61 | 1.25 | 1.4333 | 1.4 | 1.3333 | 1.2167 | 1.1667 | 2.7083 | 1.0483 | 2.9573 | 1.5333 | 3.05 | 1.8 | 2.2433 | 4.7072 |  | 3.4517
| Median |  | 3.585 | 1.25 | 1.4 | 1.4 | 1.3 | 1.2 | 1.15 | 2.705 | 1.05 | 2.9566 | 1.5 | 3.05 | 1.8 | 2.195 | 4.7234 |  | 3.405
| Std Dev. |  | 0.070427 | 0.054772 | 0.05164 | 0.089443 | 0.10328 | 0.075277 | 0.08165 | 0.0392 | 0.054924 | 0.077101 | 0.05164 | 0.054772 | 0.063246 | 0.22033 | 0.079248 |  | 0.17612
| Rel.Std.Dev. |  | 1.951% | 4.382% | 3.603% | 6.389% | 7.746% | 6.187% | 6.999% | 1.447% | 5.239% | 2.607% | 3.368% | 1.796% | 3.514% | 9.822% | 1.684% |  | 5.102%
| PDM3 |  | 79.82% | -37.73% | -28.6% | -30.26% | -33.58% | -39.39% | -41.88% | 34.91% | -47.78% | 47.31% | -23.62% | 51.93% | -10.34% | 11.75% | 134.5% |  | 71.94%
| Z-Score (Absolute) |  | 1.93 | 0.59 | 0.39 | 0.43 | 0.5 | 0.62 | 0.67 | 0.97 | 0.8 | 1.23 | 0.28 | 1.33 | 0 | 0.47 | 3.1 | 3.41 | 1.76
| Analytical results for Ti in OREAS 605b (Certified Value 457 ppm) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 0
| 1 | 1 | 485 | 420 | 400 | 460 | 470 | 480 | 460 | 440 | 498 | 468 | 458.5372 | 470 | 430 | 800 | 500 | 467 | 500 | 490.8179 | 425.8318 | 561.6 | 440 |  |  | 1
| 1 | 2 | 473 | 420 | 400 | 460 | 460 | 470 | 460 | 440 | 501 | 474 | 461.1095 | 480 | 410 | 800 | 500 | 472 | 500 | 497.4631 | 435.9325 | 554.4 | 440 |  |  | 32
| 1 | 3 | 481 | 450 | 400 | 470 | 470 | 480 | 430 | 440 | 508 | 461 | 439.659 | 480 | 410 | 800 | 500 | 471 | 500 | 492.5243 | 435.7938 | 554.5 | 450 |  |  | 16
| 1 | 4 | 472 | 450 | 400 | 470 | 470 | 480 | 430 | 440 | 504 | 483 | 461.4588 | 480 | 420 | 900 | 400 | 473 | 500 | 493.2133 | 424.9518 | 555.1 | 440 |  |  | 457.0479
| 1 | 5 | 466 | 450 | 400 | 460 | 470 | 480 | 410 | 440 | 509 | 457 | 442.787 | 470 | 380 | 800 | 500 | 470 | 500 | 484.8379 | 435.9835 | 555.7 | 440 |  |  | 108
| 1 | 6 | 478 | 450 | 400 | 450 | 470 | 440 | 420 | 440 | 503 | 442 | 457.6844 | 470 | 420 | 800 | 500 | 486 | 500 | 484.9502 | 433.0386 | 555.1 | 450
| Mean |  | 475.8333 | 440 | 400 | 461.6667 | 468.3333 | 471.6667 | 435 | 440 | 503.8333 | 464.1667 | 453.5393 | 475 | 411.6667 | 816.6667 | 483.3333 | 473.1667 | 500 | 490.6345 | 431.922 | 556.0667 | 443.3333
| Median |  | 475.5 | 450 | 400 | 460 | 470 | 480 | 430 | 440 | 503.5 | 464.5 | 458.1108 | 475 | 415 | 800 | 500 | 471.5 | 500 | 491.6711 | 434.4162 | 555.1 | 440
| Std Dev. |  | 6.8532 | 15.4919 | 0 | 7.5277 | 4.0825 | 16.0208 | 20.7364 | 0 | 4.1673 | 14.2747 | 9.7001 | 5.4772 | 17.224 | 40.8248 | 40.8248 | 6.6156 | 0 | 4.9557 | 5.1864 | 2.7515 | 5.164
| Rel.Std.Dev. |  | 1.44% | 3.521% | 0% | 1.631% | 0.8717% | 3.397% | 4.767% | 0% | 0.8271% | 3.075% | 2.139% | 1.153% | 4.184% | 4.999% | 8.447% | 1.398% | 0% | 1.01% | 1.201% | 0.4948% | 1.165%
| PDM3 |  | 4.11% | -3.73% | -12.48% | 1.011% | 2.469% | 3.199% | -4.824% | -3.73% | 10.24% | 1.558% | -0.7677% | 3.928% | -9.929% | 78.68% | 5.751% | 3.527% | 9.398% | 7.349% | -5.497% | 21.66% | -3.001%
| Z-Score (Absolute) |  | 0.27 | 0.75 | 1.89 | 0.13 | 0.06 | 0.15 | 0.89 | 0.75 | 1.07 | 0.06 | 0.36 | 0.25 | 1.56 | 9.99 | 0.49 | 0.2 | NA | 0.7 | 0.98 | 2.56 | 0.65
Results from laboratory 22 were removed due to their 100 ppm reading resolution.
| Analytical results for Tl in OREAS 605b (Certified Value 16.2 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | 25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | 01g | [5 empty] | 2
| 1 | 1 | 17.19 | 16.36 | < 10 | 14.7 | 16.15 | 14.85 | 16.75 | 15.55 | 12.4 | 15.77 | 16.2064 | 14.88 | 16.4 | 16.63 | 17.7 | 16.471 | 15.0758 | < 5 | [5 empty] | 1
| 1 | 2 | 16.9 | 16.37 | < 10 | 14.6 | 15.5 | 14.65 | 16.4 | 15.25 | 12.8 | 16.81 | 16.2952 | 15.22 | 16 | 16.51 | 17.6 | 17.112 | 16.1273 | < 5 | [5 empty] | 33
| 1 | 3 | 16.98 | 16.69 | < 10 | 15.05 | 16.1 | 16.15 | 15.6 | 15.3 | 12.6 | 16.49 | 16.0835 | 15.76 | 16.3 | 16.67 | 17.9 | 17.183 | 14.2992 | < 5 | [5 empty] | 16
| 1 | 4 | 16.73 | 16.92 | < 10 | 14.9 | 15.65 | 15.15 | 14.25 | 15.35 | 12.6 | 17.96 | 16.331 | 15.15 | 16.3 | 16.84 | 17.5 | 17.324 | 12.5856 | < 5 | [5 empty] | 16.1566
| 1 | 5 | 16.83 | 16.83 | < 10 | 14.7 | 15.45 | 14.95 | 15.95 | 15.6 | 12.7 | 16.57 | 16.1222 | 15.76 | 16.1 | 16.51 | 17.4 | 17.213 | 16.2468 | < 5 | [5 empty] | 109
| 1 | 6 | 16.7 | 16.31 | < 10 | 14.4 | 16.3 | 15.4 | 16.15 | 14.2 | 12.6 | 15.79 | 16.4472 | 16.45 | 16.7 | 16.54 | 18.1 | 16.669 | 16.0753 | < 5
| Mean |  | 16.8883 | 16.58 |  | 14.725 | 15.8583 | 15.1917 | 15.85 | 15.2083 | 12.6167 | 16.565 | 16.2476 | 15.5367 | 16.3 | 16.6167 | 17.7 | 16.9953 | 15.0683
| Median |  | 16.865 | 16.53 |  | 14.7 | 15.875 | 15.05 | 16.05 | 15.325 | 12.6 | 16.53 | 16.2508 | 15.49 | 16.3 | 16.585 | 17.65 | 17.1475 | 15.5755
| Std Dev. |  | 0.18082 | 0.26668 |  | 0.22749 | 0.36799 | 0.53518 | 0.87579 | 0.51324 | 0.13292 | 0.80552 | 0.13667 | 0.56832 | 0.24495 | 0.1277 | 0.26077 | 0.34224 | 1.4338
| Rel.Std.Dev. |  | 1.071% | 1.608% |  | 1.545% | 2.32% | 3.523% | 5.525% | 3.375% | 1.053% | 4.863% | 0.8412% | 3.658% | 1.503% | 0.7685% | 1.473% | 2.014% | 9.516%
| PDM3 |  | 4.529% | 2.62% |  | -8.861% | -1.846% | -5.973% | -1.898% | -5.869% | -21.91% | 2.528% | 0.5629% | -3.837% | 0.8874% | 2.847% | 9.553% | 5.191% | -6.736%
| Z-Score (Absolute) |  | 0.94 | 0.66 | 9.84 | 1.02 | 0 | 0.6 | 0 | 0.59 | 2.93 | 0.64 | 0.36 | 0.29 | 0.4 | 0.69 | 1.67 | 1.03 | 0.71 | 12.1
| Analytical results for Tm in OREAS 605b (Indicative Value 35.8 ppb) | [24 empty] | Indicative
Tm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab | [5 empty] | 1
| No | No | 02: 03; 03: 10; 04: 23 | [5 empty] | ppb
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS | [5 empty] | 1
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g | [5 empty] | 1
| 1 | 1 | 02: 40; 03: 40; 04: < 100 | [5 empty] | 1
| 1 | 2 | 02: 30; 03: 40; 04: < 100 | [5 empty] | 7
| 1 | 3 | 02: 30; 03: 40; 04: < 100 | [5 empty] | 16
| 1 | 4 | 02: 40; 03: 40; 04: < 100 | [5 empty] | 35.8333
| 1 | 5 | 02: < 20; 03: 40; 04: < 100 | [5 empty] | 22
| 1 | 6 | 02: 40; 03: 40; 04: < 100
| Mean |  | 02: 36; 03: 40
| Median |  | 02: 40; 03: 40
| Std Dev. |  | 02: 5.4772; 03: 0
| Rel.Std.Dev. |  | 02: 15.21%; 03: 0%
| PDM3 |  | 02: 0.4651%; 03: 11.63%
| Z-Score (Absolute) |  | 02: 0.67; 03: 0; 04: 0.81
| Analytical results for U in OREAS 605b (Certified Value 1.45 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 20 | 21 | 23 | 24 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [7 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.2g | [7 empty] | 3
| 1 | 1 | 1.59 | 1.5 | 1.4 | 1.41 | 1.48 | 1.46 | 1.52 | 1.32 | 1.58 | 1.507 | 1.3 | 1.36 | 1.36 | 1.47 | 1.448 | < 10 | [7 empty] | 1
| 1 | 2 | 1.56 | 1.5 | 1.36 | 1.43 | 1.4 | 1.46 | 1.53 | 1.38 | 1.68 | 1.5471 | 1.31 | 1.3 | 1.39 | 1.46 | 1.503 | < 10 | [7 empty] | 35
| 1 | 3 | 1.63 | 1.5 | 1.4 | 1.42 | 1.48 | 1.48 | 1.52 | 1.35 | 1.61 | 1.5164 | 1.25 | 1.32 | 1.37 | 1.51 | 1.513 | < 10 | [7 empty] | 16
| 1 | 4 | 1.55 | 1.6 | 1.38 | 1.52 | 1.47 | 1.36 | 1.55 | 1.34 | 1.65 | 1.5446 | 1.3 | 1.31 | 1.39 | 1.45 | 1.537 | < 10 | [7 empty] | 1.4532
| 1 | 5 | 1.59 | 1.4 | 1.37 | 1.43 | 1.48 | 1.52 | 1.54 | 1.36 | 1.62 | 1.5021 | 1.29 | 1.28 | 1.4 | 1.42 | 1.501 | < 10 | [7 empty] | 110
| 1 | 6 | 1.59 | 1.5 | 1.36 | 1.53 | 1.5 | 1.57 | 1.47 | 1.34 | 1.55 | 1.5088 | 1.36 | 1.34 | 1.39 | 1.48 | 1.458 | < 10
| Mean |  | 1.585 | 1.5 | 1.3783 | 1.4567 | 1.4683 | 1.475 | 1.5217 | 1.3483 | 1.615 | 1.521 | 1.3017 | 1.3183 | 1.3833 | 1.465 | 1.4933
| Median |  | 1.59 | 1.5 | 1.375 | 1.43 | 1.48 | 1.47 | 1.525 | 1.345 | 1.615 | 1.5126 | 1.3 | 1.315 | 1.39 | 1.465 | 1.502
| Std Dev. |  | 0.028107 | 0.063246 | 0.018348 | 0.053541 | 0.034881 | 0.070356 | 0.027869 | 0.020412 | 0.046797 | 0.019808 | 0.035449 | 0.028577 | 0.015055 | 0.030166 | 0.03391
| Rel.Std.Dev. |  | 1.773% | 4.216% | 1.331% | 3.676% | 2.376% | 4.77% | 1.831% | 1.514% | 2.898% | 1.302% | 2.723% | 2.168% | 1.088% | 2.059% | 2.271%
| PDM3 |  | 9.07% | 3.221% | -5.151% | 0.2392% | 1.042% | 1.501% | 4.712% | -7.216% | 11.13% | 4.667% | -10.43% | -9.28% | -4.807% | 0.8126% | 2.762%
| Z-Score (Absolute) |  | 0.93 | NA | 0.71 | 0.09 | 0 | 0.05 | 0.42 | 0.95 | 1.16 | 0.42 | 1.32 | 1.19 | 0.67 | 0.03 | 0.2 | 28.02
Results from laboratory 3 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 605b (Certified Value 24.4 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  | ppm
|  |  | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  | 2
| 1 | 1 | 26.7941 | 26 | 21 | 23 | 24 | 23 | 25 | 25 | 23 | 25 | 27 | 23.7388 | 24 | 22.5 | 24 | 27 | 23.4 | 19 | 26.2091 | 24.2378 | 30.47 | 22.732 |  | 1
| 1 | 2 | 25.5729 | 27 | 22 | 24 | 24 | 23 | 25 | 24 | 23 | 25 | 28 | 24.3419 | 24 | 21.5 | 25 | 26 | 23.7 | 19 | 26.3067 | 24.3978 | 30.38 | 23.629 |  | 36
| 1 | 3 | 26.4694 | 26 | 22 | 23 | 25 | 23 | 25 | 22 | 23 | 26 | 27 | 23.4231 | 24 | 22.1 | 24 | 26 | 23 | 19 | 25.8163 | 24.766 | 29.65 | 23.227 |  | 16
| 1 | 4 | 26.8394 | 25 | 23 | 24 | 25 | 24 | 25 | 22 | 24 | 26 | 28 | 25.1468 | 25 | 21.6 | 26 | 26 | 23.5 | 20 | 26.2951 | 24.1683 | 30.55 | 24.298 |  | 24.4072
| 1 | 5 | 25.4939 | 26 | 23 | 23 | 25 | 23 | 25 | 21 | 24 | 26 | 26 | 24.3032 | 24 | 22.6 | 26 | 26 | 22.9 | 19 | 25.5895 | 24.4434 | 29.52 | 23.428 |  | 111
| 1 | 6 | 26.2139 | 27 | 23 | 24 | 24 | 23 | 23 | 21 | 24 | 26 | 26 | 24.5074 | 24 | 21.4 | 25 | 26 | 23.9 | 18 | 25.4657 | 24.475 | 29.45 | 24.429
| Mean |  | 26.2306 | 26.1667 | 22.3333 | 23.5 | 24.5 | 23.1667 | 24.6667 | 22.5 | 23.5 | 25.6667 | 27 | 24.2435 | 24.1667 | 21.95 | 25 | 26.1667 | 23.4 | 19 | 25.9471 | 24.4147 | 30.0033 | 23.6238
| Median |  | 26.3416 | 26 | 22.5 | 23.5 | 24.5 | 23 | 25 | 22 | 23.5 | 26 | 27 | 24.3226 | 24 | 21.85 | 25 | 26 | 23.45 | 19 | 26.0127 | 24.4206 | 30.015 | 23.5285
| Std Dev. |  | 0.58666 | 0.75277 | 0.8165 | 0.54772 | 0.54772 | 0.40825 | 0.8165 | 1.6432 | 0.54772 | 0.5164 | 0.89443 | 0.60441 | 0.40825 | 0.5244 | 0.89443 | 0.40825 | 0.38987 | 0.63246 | 0.37304 | 0.20985 | 0.51442 | 0.64715
| Rel.Std.Dev. |  | 2.237% | 2.877% | 3.656% | 2.331% | 2.236% | 1.762% | 3.31% | 7.303% | 2.331% | 2.012% | 3.313% | 2.493% | 1.689% | 2.389% | 3.578% | 1.56% | 1.666% | 3.329% | 1.438% | 0.8595% | 1.715% | 2.739%
| PDM3 |  | 7.471% | 7.209% | -8.497% | -3.717% | 0.3804% | -5.082% | 1.063% | -7.814% | -3.717% | 5.16% | 10.62% | -0.6704% | -0.9853% | -10.07% | 2.429% | 7.209% | -4.126% | -22.15% | 6.309% | 0.03103% | 22.93% | -3.209%
| Z-Score (Absolute) |  | 1.03 | 0.99 | 1.08 | 0.45 | 0.09 | 0.63 | 0.18 | 0.99 | 0.45 | 0.72 | 1.44 | 0.05 | 0.09 | 1.28 | 0.36 | 0.99 | 0.5 | 2.87 | 0.87 | 0.05 | 3.06 | 0.38
| Analytical results for W in OREAS 605b (Certified Value 8 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.25g | 0.2g | 01g | 0.25g |  |  | 2
| 1 | 1 | 9.1 | 5.6 | < 5 | 8.02 | 8.09 | 7.92 | 7.75 | 7.01 | 6.02 | 8.98 | 8.7046 | 7.15 | 9.44 | < 10 | 8.2 | 9 | 15 | 8.7703 | < 20 | < 5 | < 10 |  |  | 1
| 1 | 2 | 9.1 | 5.4 | < 5 | 8.08 | 7.86 | 7.47 | 7.89 | 7.35 | 6.28 | 9.43 | 8.7077 | 6.62 | 9.45 | < 10 | 8.3 | 8.95 | 14 | 9.0761 | < 20 | < 5 | < 10 |  |  | 37
| 1 | 3 | 8.8 | 5.9 | < 5 | 8.3 | 8.48 | 8.05 | 7.75 | 7.01 | 6.24 | 9.35 | 8.2639 | 7.06 | 9.44 | < 10 | 8.3 | 9.02 | 19 | 9.064 | < 20 | < 5 | < 10 |  |  | 16
| 1 | 4 | 8.8 | 6.1 | < 5 | 8.2 | 8.09 | 7.92 | 7.02 | 7.94 | 6.1 | 9.18 | 8.5396 | 6.79 | 10.2 | < 10 | 8.4 | 8.76 | 17 | 9.0013 | < 20 | < 5 | < 10 |  |  | 8.0011
| 1 | 5 | 8.3 | 5.7 | < 5 | 8.16 | 8.01 | 7.77 | 7.53 | 7.16 | 6.17 | 9.36 | 8.0855 | 7.31 | 9.38 | < 10 | 8.2 | 8.93 | 17 | 9.3016 | < 20 | < 5 | < 10 |  |  | 112
| 1 | 6 | 9.3 | 5.6 | < 5 | 8.1 | 8.49 | 8.11 | 7.4 | 7.14 | 6.11 | 8.84 | 8.2148 | 7.32 | 9.61 | < 10 | 8.2 | 9.13 | 16 | 8.921 | < 20 | < 5 | < 10
| Mean |  | 8.9 | 5.7167 |  | 8.1433 | 8.17 | 7.8733 | 7.5567 | 7.2683 | 6.1533 | 9.19 | 8.4193 | 7.0417 | 9.5867 |  | 8.2667 | 8.965 | 16.3333 | 9.0224
| Median |  | 8.95 | 5.65 |  | 8.13 | 8.09 | 7.92 | 7.64 | 7.15 | 6.14 | 9.265 | 8.4017 | 7.105 | 9.445 |  | 8.25 | 8.975 | 16.5 | 9.0327
| Std Dev. |  | 0.35214 | 0.24833 |  | 0.09913 | 0.25807 | 0.23019 | 0.31595 | 0.35199 | 0.096264 | 0.23614 | 0.267 | 0.28379 | 0.3102 |  | 0.08165 | 0.12243 | 1.7512 | 0.17711
| Rel.Std.Dev. |  | 3.957% | 4.344% |  | 1.217% | 3.159% | 2.924% | 4.181% | 4.843% | 1.564% | 2.569% | 3.171% | 4.03% | 3.236% |  | 0.9877% | 1.366% | 10.72% | 1.963%
| PDM3 |  | 11.23% | -28.55% |  | 1.778% | 2.111% | -1.597% | -5.555% | -9.158% | -23.09% | 14.86% | 5.227% | -11.99% | 19.82% |  | 3.319% | 12.05% | 104.1% | 12.76%
| Z-Score (Absolute) |  | 0.49 | 1.56 | 3.64 | 0 | 0.02 | 0.17 | 0.38 | 0.56 | 1.28 | 0.67 | 0.18 | 0.71 | 0.93 | 2.03 | 0.08 | 0.53 | 5.28 | 0.57 | 1.2 | 3.64 | 2.03
| Analytical results for Y in OREAS 605b (Certified Value 3.37 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 26 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 0.25g |  |  |  | 3
| 1 | 1 | 3.45 | 3.23 | 9 | 3.35 | 3.21 | 3.79 | 3.47 | 3.58 | 3.26 | 3.57 | 3.4606 | 3.14 | 2.95 | 3.3 | 3.18 | 3.37 | < 5 | 3.1941 | 3.2006 | 3.279 |  |  |  | 1
| 1 | 2 | 3.36 | 3.16 | 9 | 3.41 | 3.27 | 3.6 | 3.46 | 3.56 | 3.34 | 3.67 | 3.5757 | 3.1 | 2.82 | 3.2 | 3.22 | 3.37 | < 5 | 3.1888 | 3.2179 | 3.406 |  |  |  | 38
| 1 | 3 | 3.37 | 3.4 | 9 | 3.42 | 3.33 | 3.71 | 3.64 | 3.45 | 3.26 | 3.63 | 3.4076 | 3.4 | 2.89 | 3.2 | 3.23 | 3.44 | < 5 | 3.0975 | 3.289 | 3.288 |  |  |  | 16
| 1 | 4 | 3.38 | 3.37 | 9 | 3.29 | 3.32 | 3.53 | 3.35 | 3.52 | 3.23 | 3.57 | 3.5526 | 2.96 | 2.88 | 3.4 | 3.23 | 3.39 | < 5 | 3.1888 | 3.1981 | 3.441 |  |  |  | 3.3679
| 1 | 5 | 3.33 | 3.38 | 9 | 3.39 | 3.23 | 3.67 | 3.77 | 3.4 | 3.32 | 3.61 | 3.4224 | 3.4 | 2.8 | 3.5 | 3.19 | 3.39 | < 5 | 3.1521 | 3.2382 | 3.305 |  |  |  | 113
| 1 | 6 | 3.54 | 3.32 | 9 | 3.42 | 3.35 | 3.4 | 3.73 | 3.37 | 3.31 | 3.51 | 3.5221 | 3.04 | 2.86 | 3.4 | 3.21 | 3.5 | < 5 | 3.0859 | 3.2989 | 3.464
| Mean |  | 3.405 | 3.31 | 9 | 3.38 | 3.285 | 3.6167 | 3.57 | 3.48 | 3.2867 | 3.5933 | 3.4902 | 3.1733 | 2.8667 | 3.3333 | 3.21 | 3.41 |  | 3.1512 | 3.2404 | 3.3638
| Median |  | 3.375 | 3.345 | 9 | 3.4 | 3.295 | 3.635 | 3.555 | 3.485 | 3.285 | 3.59 | 3.4914 | 3.12 | 2.87 | 3.35 | 3.215 | 3.39 |  | 3.1705 | 3.228 | 3.3555
| Std Dev. |  | 0.077136 | 0.095499 | 0 | 0.051381 | 0.057184 | 0.1388 | 0.16793 | 0.086487 | 0.04274 | 0.055737 | 0.070037 | 0.18576 | 0.053541 | 0.12111 | 0.020976 | 0.05099 |  | 0.048616 | 0.043967 | 0.082674
| Rel.Std.Dev. |  | 2.265% | 2.885% | 0% | 1.52% | 1.741% | 3.838% | 4.704% | 2.485% | 1.3% | 1.551% | 2.007% | 5.854% | 1.868% | 3.633% | 0.6535% | 1.495% |  | 1.543% | 1.357% | 2.458%
| PDM3 |  | 1.102% | -1.719% | 167.2% | 0.3598% | -2.461% | 7.387% | 6.001% | 3.329% | -2.411% | 6.694% | 3.631% | -5.777% | -14.88% | -1.026% | -4.688% | 1.251% |  | -6.433% | -3.784% | -0.1202%
| Z-Score (Absolute) |  | 0.39 | 0.13 | NA | 0.26 | 0.26 | 1.55 | 1.29 | 0.8 | 0.26 | 1.42 | 0.86 | 0.87 | 2.55 | 0 | 0.67 | 0.42 | 4.56 | 1 | 0.51 | 0.17
Results from laboratory 4 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 605b (Certified Value 0.26 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  | 1
| No | No | 02: 03; 03: 10; 04: 20; 05: 21; 06: 23 |  |  |  | ppm
|  |  | 02: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS |  |  |  | 2
|  |  | 02: 0.5g; 03: 01g; 04: 0.25g; 05: 0.25g; 06: 0.25g |  |  |  | 3
| 1 | 1 | 02: 0.24; 03: 0.29; 04: 0.3; 05: 0.262; 06: 0.247 |  |  |  | 1
| 1 | 2 | 02: 0.24; 03: 0.31; 04: 0.3; 05: 0.246; 06: 0.25 |  |  |  | 39
| 1 | 3 | 02: 0.26; 03: 0.29; 04: 0.3; 05: 0.254; 06: 0.241 |  |  |  | 16
| 1 | 4 | 02: 0.27; 03: 0.27; 04: 0.3; 05: 0.253; 06: 0.254 |  |  |  | 0.26493
| 1 | 5 | 02: 0.25; 03: 0.28; 04: 0.3; 05: 0.253; 06: 0.241 |  |  |  | 114
| 1 | 6 | 02: 0.28; 03: 0.28; 04: 0.2; 05: 0.247; 06: 0.24
| Mean |  | 02: 0.25667; 03: 0.28667; 04: 0.28333; 05: 0.2525; 06: 0.2455
| Median |  | 02: 0.255; 03: 0.285; 04: 0.3; 05: 0.253; 06: 0.244
| Std Dev. |  | 02: 0.01633; 03: 0.013663; 04: 0.040825; 05: 0.0057533; 06: 0.0057533
| Rel.Std.Dev. |  | 02: 6.362%; 03: 4.766%; 04: 14.41%; 05: 2.279%; 06: 2.343%
| PDM3 |  | 02: -3.12%; 03: 8.203%; 04: 6.945%; 05: -4.693%; 06: -7.335%
| Z-Score (Absolute) |  | 02: 0; 03: 1.81; 04: 1.61; 05: 0.25; 06: 0.67
| Analytical results for Zn in OREAS 605b (Certified Value 0.232 wt.%) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/AAS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.2252 | 0.23672 | 0.216 | 0.221 | 0.23 | 0.241 | 0.243 | 0.223 | 0.228 | 0.2386 | 0.22227 | 0.23421 | 0.245 | 0.238 | 0.21 | 0.227 | 0.237 | 0.24839 | 0.23335 | 0.2309 | 0.24475 |  |  | 1
| 1 | 2 | 0.2232 | 0.2364 | 0.216 | 0.22 | 0.23 | 0.237 | 0.244 | 0.222 | 0.228 | 0.2459 | 0.22927 | 0.23607 | 0.236 | 0.235 | 0.2103 | 0.23 | 0.229 | 0.25272 | 0.23168 | 0.2337 | 0.24464 |  |  | 40
| 1 | 3 | 0.2275 | 0.24876 | 0.216 | 0.224 | 0.234 | 0.241 | 0.24 | 0.223 | 0.228 | 0.2421 | 0.22276 | 0.2321 | 0.232 | 0.234 | 0.2153 | 0.236 | 0.228 | 0.24102 | 0.23621 | 0.2298 | 0.24253 |  |  | 16
| 1 | 4 | 0.2239 | 0.24993 | 0.216 | 0.224 | 0.232 | 0.24 | 0.243 | 0.23 | 0.229 | 0.2429 | 0.23257 | 0.23902 | 0.24 | 0.251 | 0.2137 | 0.236 | 0.231 | 0.24738 | 0.23054 | 0.2313 | 0.24322 |  |  | 0.23225
| 1 | 5 | 0.2182 | 0.2484 | 0.215 | 0.221 | 0.234 | 0.236 | 0.237 | 0.226 | 0.227 | 0.2462 | 0.22319 | 0.22949 | 0.23 | 0.241 | 0.2085 | 0.232 | 0.229 | 0.24242 | 0.23478 | 0.2311 | 0.24256 |  |  | 115
| 1 | 6 | 0.2214 | 0.24556 | 0.216 | 0.215 | 0.236 | 0.22 | 0.242 | 0.226 | 0.228 | 0.2323 | 0.22811 | 0.23506 | 0.236 | 0.232 | 0.21 | 0.238 | 0.233 | 0.24163 | 0.23368 | 0.2298 | 0.24694
| Mean |  | 0.22323 | 0.2443 | 0.21583 | 0.22083 | 0.23267 | 0.23583 | 0.2415 | 0.225 | 0.228 | 0.24133 | 0.22636 | 0.23433 | 0.2365 | 0.2385 | 0.2113 | 0.23317 | 0.23117 | 0.24559 | 0.23337 | 0.2311 | 0.24411
| Median |  | 0.22355 | 0.24698 | 0.216 | 0.221 | 0.233 | 0.2385 | 0.2425 | 0.2245 | 0.228 | 0.2425 | 0.22565 | 0.23464 | 0.236 | 0.2365 | 0.21015 | 0.234 | 0.23 | 0.2449 | 0.23351 | 0.231 | 0.24393
| Std Dev. |  | 0.0032004 | 0.006162 | 0.00040825 | 0.0033116 | 0.0024221 | 0.0080353 | 0.0025884 | 0.0029665 | 0.00063246 | 0.0052279 | 0.0042393 | 0.0032854 | 0.0054314 | 0.006892 | 0.0026069 | 0.0042151 | 0.0033714 | 0.0046573 | 0.0020496 | 0.0014297 | 0.0016974
| Rel.Std.Dev. |  | 1.434% | 2.522% | 0.1891% | 1.5% | 1.041% | 3.407% | 1.072% | 1.318% | 0.2774% | 2.166% | 1.873% | 1.402% | 2.297% | 2.89% | 1.234% | 1.808% | 1.458% | 1.896% | 0.8782% | 0.6186% | 0.6954%
| PDM3 |  | -3.881% | 5.188% | -7.067% | -4.914% | 0.1807% | 1.544% | 3.984% | -3.12% | -1.829% | 3.912% | -2.534% | 0.8947% | 1.831% | 2.692% | -9.019% | 0.396% | -0.4652% | 5.747% | 0.4846% | -0.4939% | 5.106%
| Z-Score (Absolute) |  | 0.98 | 1.1 | 1.72 | 1.22 | 0.05 | 0.26 | 0.83 | 0.81 | 0.51 | 0.81 | 0.67 | 0.11 | 0.33 | 0.53 | 2.17 | 0 | 0.2 | 1.23 | 0.02 | 0.2 | 1.08
| Analytical results for Zr in OREAS 605b (Certified Value 13.3 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 1
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 10g | 01g | 0.5g | 0.5g | 10g | 0.25g | 0.25g | 10g | 0.5g | 0.2g | 01g | 0.25g |  |  | 1
| 1 | 1 | 15.7 | 12.6 | 15 | 14.1 | 12.9 | 15.5 | 13.2 | 13.8 | 13.1 | 15.3 | 13.9301 | 12.4 | 11.3 | 18.1 | 11.4 | 14.6 | 12.9 | 12.058 | 11.6954 | 16.21 | 11.16 |  |  | 1
| 1 | 2 | 15.4 | 12.7 | 14 | 14.5 | 12.5 | 14.5 | 13.1 | 13.4 | 13 | 15.6 | 14.4247 | 11.4 | 11.3 | 18.2 | 11.4 | 15 | 12.6 | 12.204 | 12.0189 | 15.89 | 11.18 |  |  | 41
| 1 | 3 | 15.3 | 13.3 | 14 | 14.4 | 13.1 | 15.3 | 12.6 | 12.9 | 13 | 15.6 | 13.7907 | 12.4 | 10.3 | 17.6 | 11.8 | 14.4 | 12.9 | 11.86 | 12.2031 | 15.71 | 11.51 |  |  | 16
| 1 | 4 | 15.4 | 13.6 | 14 | 14 | 12.9 | 14.5 | 11.6 | 12.8 | 13 | 16.3 | 14.4014 | 12 | 10.6 | 19.4 | 11.7 | 14.7 | 13 | 12.018 | 12.0881 | 16.72 | 11.69 |  |  | 13.2556
| 1 | 5 | 14.7 | 13 | 14 | 14.2 | 13 | 14.7 | 13.2 | 13.5 | 12.9 | 15.3 | 13.8423 | 12.3 | 10.7 | 19.5 | 12.7 | 14.2 | 13 | 11.642 | 11.6598 | 14.39 | 11.48 |  |  | 116
| 1 | 6 | 16 | 13.2 | 15 | 13.7 | 13 | 13.5 | 13.2 | 12.8 | 13 | 15.2 | 13.9916 | 10.8 | 10.5 | 19.4 | 11.4 | 14.8 | 12.3 | 11.799 | 11.9059 | 15.38 | 11.52
| Mean |  | 15.4167 | 13.0667 | 14.3333 | 14.15 | 12.9 | 14.6667 | 12.8167 | 13.2 | 13 | 15.55 | 14.0635 | 11.8833 | 10.7833 | 18.7 | 11.7333 | 14.6167 | 12.7833 | 11.9302 | 11.9285 | 15.7167 | 11.4233
| Median |  | 15.4 | 13.1 | 14 | 14.15 | 12.95 | 14.6 | 13.15 | 13.15 | 13 | 15.45 | 13.9608 | 12.15 | 10.65 | 18.8 | 11.55 | 14.65 | 12.9 | 11.939 | 11.9624 | 15.8 | 11.495
| Std Dev. |  | 0.43551 | 0.37771 | 0.5164 | 0.2881 | 0.20976 | 0.70899 | 0.64005 | 0.42426 | 0.063246 | 0.40373 | 0.27964 | 0.65243 | 0.42151 | 0.82946 | 0.50465 | 0.28577 | 0.27869 | 0.20209 | 0.21732 | 0.79417 | 0.20963
| Rel.Std.Dev. |  | 2.825% | 2.891% | 3.603% | 2.036% | 1.626% | 4.834% | 4.994% | 3.214% | 0.4865% | 2.596% | 1.988% | 5.49% | 3.909% | 4.436% | 4.301% | 1.955% | 2.18% | 1.694% | 1.822% | 5.053% | 1.835%
| PDM3 |  | 16.3% | -1.425% | 8.131% | 6.748% | -2.682% | 10.65% | -3.311% | -0.4191% | -1.928% | 17.31% | 6.095% | -10.35% | -18.65% | 41.07% | -11.48% | 10.27% | -3.562% | -9.999% | -10.01% | 18.57% | -13.82%
| Z-Score (Absolute) |  | 1.42 | 0.02 | NA | 0.66 | 0.08 | 0.97 | 0.13 | 0.1 | 0.02 | 1.5 | 0.61 | 0.68 | 1.34 | 3.37 | 0.77 | 0.94 | 0.15 | 0.66 | 0.66 | 1.6 | 0.96
Results from laboratory 4 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 605b (Indicative Value 12.1 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.09 | 1
| 1 | 2 | 12.11 | 18
| Mean |  | 12.1 | 16
| Median |  | 12.1 | 12.1
| Std Dev. |  | 0.014142 | 24
| Rel.Std.Dev. |  | 0.1169%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 605b (Indicative Value 0.595 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.6 | 19
| Mean |  | 0.595 | 16
| Median |  | 0.595 | 0.595
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 1.188%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 605b (Indicative Value 10.27 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 10.26 | 1
| 1 | 2 | 10.27 | 20
| Mean |  | 10.265 | 16
| Median |  | 10.265 | 10.265
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 0.06889%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 605b (Indicative Value 1.46 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.45 | 1
| 1 | 2 | 1.46 | 21
| Mean |  | 1.455 | 16
| Median |  | 1.455 | 1.455
| Std Dev. |  | 0.0070711 | 27
| Rel.Std.Dev. |  | 0.486%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 605b (Indicative Value 0.345 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.35 | 18
| Mean |  | 0.345 | 16
| Median |  | 0.345 | 0.345
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 605b (Indicative Value 0.02 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.02 | 19
| Mean |  | 0.02 | 16
| Median |  | 0.02 | 0.02
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 605b (Indicative Value 0.81 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.81 | 1
| 1 | 2 | 0.81 | 20
| Mean |  | 0.81 | 16
| Median |  | 0.81 | 0.81
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 605b (Indicative Value 0.157 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.157 | 1
| 1 | 2 | 0.156 | 21
| Mean |  | 0.1565 | 16
| Median |  | 0.1565 | 0.1565
| Std Dev. |  | 0.00070711 | 27
| Rel.Std.Dev. |  | 0.4518%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 605b (Indicative Value 51.9 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 51.92 | 1
| 1 | 2 | 51.88 | 18
| Mean |  | 51.9 | 16
| Median |  | 51.9 | 51.9
| Std Dev. |  | 0.028284 | 24
| Rel.Std.Dev. |  | 0.0545%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 605b (Indicative Value 26.05 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 26.1 | 1
| 1 | 2 | 26 | 19
| Mean |  | 26.05 | 16
| Median |  | 26.05 | 26.05
| Std Dev. |  | 0.070711 | 25
| Rel.Std.Dev. |  | 0.2714%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 605b (Indicative Value 0.345 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.35 | 20
| Mean |  | 0.345 | 16
| Median |  | 0.345 | 0.345
| Std Dev. |  | 0.0070711 | 26
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 605b (Indicative Value 1.29 wt.%) |  |  |  | Indicative
H2O- | Round | Replicate | Lab |  | 1
| No | No | 23 |  | wt.%
|  |  | at 105°C*Oven |  | 2
2
| 1 | 1 | 1.54 |  | 1
| 1 | 2 | 1.36 |  | 23
| 1 | 3 | 1.4 |  | 16
| 1 | 4 | 1.09 |  | 1.2883
| 1 | 5 | 1.2 |  | 29
| 1 | 6 | 1.14
| Mean |  | 1.2883
| Median |  | 1.28
| Std Dev. |  | 0.17325
| Rel.Std.Dev. |  | 13.45%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI1000 in OREAS 605b (Indicative Value 13.32 wt.%) |  |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | Lab | 1
| No | No | Perth | 23 | wt.%
|  |  | LOI*TGA | LOI*TGA | 2
2
| 1 | 1 |  | 13.27 | 1
| 1 | 2 | 13.09 | 13.32 | 23
| 1 | 3 |  | 13.42 | 16
| 1 | 4 | 13.11 | 13.35 | 13.3192
| 1 | 5 |  | 13.92 | 29
| 1 | 6 |  | 13.95
| Mean |  | 13.1 | 13.5383
| Median |  | 13.1 | 13.385
| Std Dev. |  | 0.014142 | 0.31122
| Rel.Std.Dev. |  | 0.108% | 2.299%
| PDM3 |  | -1.645% | 1.645%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 605b (Indicative Value 1045 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1050 | 1
| 1 | 2 | 1040 | 25
| Mean |  | 1045 | 16
| Median |  | 1045 | 1045
| Std Dev. |  | 7.0711 | 31
| Rel.Std.Dev. |  | 0.6767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 605b (Indicative Value 2480 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2530 | 1
| 1 | 2 | 2430 | 26
| Mean |  | 2480 | 16
| Median |  | 2480 | 2480
| Std Dev. |  | 70.7107 | 32
| Rel.Std.Dev. |  | 2.851%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 605b (Indicative Value 12950 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 12900 | 1
| 1 | 2 | 13000 | 6
| Mean |  | 12950 | 16
| Median |  | 12950 | 12950
| Std Dev. |  | 70.7107 | 33
| Rel.Std.Dev. |  | 0.546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 605b (Indicative Value 0.7 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.8 | 28
| Mean |  | 0.7 | 16
| Median |  | 0.7 | 0.7
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 605b (Indicative Value 54 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 54 | 1
| 1 | 2 | 53.5 | 29
| Mean |  | 53.75 | 16
| Median |  | 53.75 | 53.75
| Std Dev. |  | 0.35355 | 35
| Rel.Std.Dev. |  | 0.6578%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 605b (Indicative Value 17.5 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.4 | 1
| 1 | 2 | 17.6 | 30
| Mean |  | 17.5 | 16
| Median |  | 17.5 | 17.5
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 605b (Indicative Value 32.3 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32.3 | 1
| 1 | 2 | 32.2 | 31
| Mean |  | 32.25 | 16
| Median |  | 32.25 | 32.25
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.2193%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 605b (Indicative Value 19 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.4 | 1
| 1 | 2 | 18.6 | 32
| Mean |  | 19 | 16
| Median |  | 19 | 19
| Std Dev. |  | 0.56569 | 38
| Rel.Std.Dev. |  | 2.977%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 605b (Indicative Value 35 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38 | 1
| 1 | 2 | 32 | 33
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 4.2426 | 39
| Rel.Std.Dev. |  | 12.12%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 605b (Indicative Value 2.4 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.39 | 1
| 1 | 2 | 2.4 | 34
| Mean |  | 2.395 | 16
| Median |  | 2.395 | 2.395
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 0.2952%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 605b (Indicative Value 4.68 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.74 | 1
| 1 | 2 | 4.62 | 35
| Mean |  | 4.68 | 16
| Median |  | 4.68 | 4.68
| Std Dev. |  | 0.084853 | 41
| Rel.Std.Dev. |  | 1.813%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 605b (Indicative Value 1.28 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.25 | 1
| 1 | 2 | 1.3 | 12
| Mean |  | 1.275 | 16
| Median |  | 1.275 | 1.275
| Std Dev. |  | 0.035355 | 42
| Rel.Std.Dev. |  | 2.773%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 605b (Indicative Value 0.61 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.62 | 13
| Mean |  | 0.605 | 16
| Median |  | 0.605 | 0.605
| Std Dev. |  | 0.021213 | 43
| Rel.Std.Dev. |  | 3.506%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 605b (Indicative Value 0.39 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.37 | 14
| Mean |  | 0.385 | 16
| Median |  | 0.385 | 0.385
| Std Dev. |  | 0.021213 | 44
| Rel.Std.Dev. |  | 5.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 605b (Indicative Value 31 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.1 | 1
| 1 | 2 | 30.9 | 39
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 0.14142 | 45
| Rel.Std.Dev. |  | 0.4562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 605b (Indicative Value 1.69 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.66 | 1
| 1 | 2 | 1.72 | 15
| Mean |  | 1.69 | 16
| Median |  | 1.69 | 1.69
| Std Dev. |  | 0.042426 | 46
| Rel.Std.Dev. |  | 2.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 605b (Indicative Value 6.65 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.5 | 1
| 1 | 2 | 6.8 | 16
| Mean |  | 6.65 | 16
| Median |  | 6.65 | 6.65
| Std Dev. |  | 0.21213 | 47
| Rel.Std.Dev. |  | 3.19%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 605b (Indicative Value 3.25 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.31 | 1
| 1 | 2 | 3.18 | 25
| Mean |  | 3.245 | 16
| Median |  | 3.245 | 3.245
| Std Dev. |  | 0.091924 | 31
| Rel.Std.Dev. |  | 2.833%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 605b (Indicative Value 0.25 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.24 | 10
| Mean |  | 0.245 | 16
| Median |  | 0.245 | 0.245
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 605b (Indicative Value 3.63 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.65 | 1
| 1 | 2 | 3.6 | 27
| Mean |  | 3.625 | 16
| Median |  | 3.625 | 3.625
| Std Dev. |  | 0.035355 | 33
| Rel.Std.Dev. |  | 0.9753%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 605b (Indicative Value 18.6 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.4 | 1
| 1 | 2 | 18.8 | 5
| Mean |  | 18.6 | 16
| Median |  | 18.6 | 18.6
| Std Dev. |  | 0.28284 | 34
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 605b (Indicative Value 90 ppb) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 90 | 1
| 1 | 2 | 90 | 29
| Mean |  | 90 | 16
| Median |  | 90 | 90
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 605b (Indicative Value 109 ppm) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 110 | 1
| 1 | 2 | 108 | 30
| Mean |  | 109 | 16
| Median |  | 109 | 109
| Std Dev. |  | 1.4142 | 36
| Rel.Std.Dev. |  | 1.297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 605b (Indicative Value 17.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.6 | 1
| 1 | 2 | 17.6 | 31
| Mean |  | 17.6 | 16
| Median |  | 17.6 | 17.6
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 605b (Indicative Value 6.92 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.87 | 1
| 1 | 2 | 6.96 | 32
| Mean |  | 6.915 | 16
| Median |  | 6.915 | 6.915
| Std Dev. |  | 0.06364 | 38
| Rel.Std.Dev. |  | 0.9203%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 605b (Indicative Value 12.5 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.2 | 1
| 1 | 2 | 12.7 | 13
| Mean |  | 12.45 | 16
| Median |  | 12.45 | 12.45
| Std Dev. |  | 0.35355 | 39
| Rel.Std.Dev. |  | 2.84%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 605b (Indicative Value 24 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26 | 1
| 1 | 2 | 22 | 34
| Mean |  | 24 | 16
| Median |  | 24 | 24
| Std Dev. |  | 2.8284 | 40
| Rel.Std.Dev. |  | 11.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 605b (Indicative Value 1560 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1570 | 1
| 1 | 2 | 1550 | 35
| Mean |  | 1560 | 16
| Median |  | 1560 | 1560
| Std Dev. |  | 14.1421 | 41
| Rel.Std.Dev. |  | 0.9065%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 605b (Indicative Value 3.54 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.54 | 1
| 1 | 2 | 3.54 | 15
| Mean |  | 3.54 | 16
| Median |  | 3.54 | 3.54
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 605b (Indicative Value 38.8 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.9 | 1
| 1 | 2 | 38.7 | 37
| Mean |  | 38.8 | 16
| Median |  | 38.8 | 38.8
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 0.3645%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 605b (Indicative Value 30 ppb) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 10 | 8
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 28.2843 | 44
| Rel.Std.Dev. |  | 94.28%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 605b (Indicative Value 457 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 463 | 1
| 1 | 2 | 451 | 39
| Mean |  | 457 | 16
| Median |  | 457 | 457
| Std Dev. |  | 8.4853 | 45
| Rel.Std.Dev. |  | 1.857%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 605b (Indicative Value 8.2 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.3 | 1
| 1 | 2 | 8.1 | 40
| Mean |  | 8.2 | 16
| Median |  | 8.2 | 8.2
| Std Dev. |  | 0.14142 | 46
| Rel.Std.Dev. |  | 1.725%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 605b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 605b (Indicative Value 2.22 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.24 | 12
| Mean |  | 2.22 | 16
| Median |  | 2.22 | 2.22
| Std Dev. |  | 0.028284 | 31
| Rel.Std.Dev. |  | 1.274%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 605b (Indicative Value 7.3 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.2 | 1
| 1 | 2 | 7.4 | 26
| Mean |  | 7.3 | 16
| Median |  | 7.3 | 7.3
| Std Dev. |  | 0.14142 | 32
| Rel.Std.Dev. |  | 1.937%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 605b (Indicative Value 558 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 560 | 1
| 1 | 2 | 555 | 27
| Mean |  | 557.5 | 16
| Median |  | 557.5 | 557.5
| Std Dev. |  | 3.5355 | 33
| Rel.Std.Dev. |  | 0.6342%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 605b (Indicative Value 0.55 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.54 | 1
| 1 | 2 | 0.56 | 13
| Mean |  | 0.55 | 16
| Median |  | 0.55 | 0.55
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 605b (Indicative Value 0.25 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.24 | 1
| 1 | 2 | 0.25 | 14
| Mean |  | 0.245 | 16
| Median |  | 0.245 | 0.245
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 605b (Indicative Value 34.6 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.8 | 1
| 1 | 2 | 33.4 | 30
| Mean |  | 34.6 | 16
| Median |  | 34.6 | 34.6
| Std Dev. |  | 1.6971 | 36
| Rel.Std.Dev. |  | 4.905%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 605b (Indicative Value 7 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.98 | 1
| 1 | 2 | 7.01 | 15
| Mean |  | 6.995 | 16
| Median |  | 6.995 | 6.995
| Std Dev. |  | 0.021213 | 37
| Rel.Std.Dev. |  | 0.3033%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 605b (Indicative Value 0.213 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.212 | 1
| 1 | 2 | 0.214 | 32
| Mean |  | 0.213 | 16
| Median |  | 0.213 | 0.213
| Std Dev. |  | 0.0014142 | 38
| Rel.Std.Dev. |  | 0.664%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 605b (Indicative Value 16.8 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17 | 1
| 1 | 2 | 16.6 | 33
| Mean |  | 16.8 | 16
| Median |  | 16.8 | 16.8
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 1.684%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 605b (Indicative Value 90 ppb) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 90 | 1
| 1 | 2 | 90 | 7
| Mean |  | 90 | 16
| Median |  | 90 | 90
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 605b (Indicative Value 3.07 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.03 | 1
| 1 | 2 | 3.11 | 35
| Mean |  | 3.07 | 16
| Median |  | 3.07 | 3.07
| Std Dev. |  | 0.056569 | 41
| Rel.Std.Dev. |  | 1.843%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 605b (Indicative Value 59 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 59.3 | 1
| 1 | 2 | 59.3 | 36
| Mean |  | 59.3 | 16
| Median |  | 59.3 | 59.3
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 605b (Indicative Value 30 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 37
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 605b (Indicative Value 6.64 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.58 | 1
| 1 | 2 | 6.7 | 38
| Mean |  | 6.64 | 16
| Median |  | 6.64 | 6.64
| Std Dev. |  | 0.084853 | 44
| Rel.Std.Dev. |  | 1.278%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 605b (Indicative Value 0.63 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.66 | 1
| 1 | 2 | 0.59 | 39
| Mean |  | 0.625 | 16
| Median |  | 0.625 | 0.625
| Std Dev. |  | 0.049497 | 45
| Rel.Std.Dev. |  | 7.92%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 605b (Indicative Value 0.231 wt.%) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.234 | 1
| 1 | 2 | 0.227 | 40
| Mean |  | 0.2305 | 16
| Median |  | 0.2305 | 0.2305
| Std Dev. |  | 0.0049497 | 46
| Rel.Std.Dev. |  | 2.147%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 605b (Indicative Value 120 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 119 | 1
| 1 | 2 | 120 | 41
| Mean |  | 119.5 | 16
| Median |  | 119.5 | 119.5
| Std Dev. |  | 0.70711 | 47
| Rel.Std.Dev. |  | 0.5917%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.