## Abbreviations
Table 1. Abbreviations used for OREAS 461
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 461
Lab Slot | Laboratory Name and Location
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
Nagrom, Perth, WA, Australia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
SGS Vostok Limited, Chita, Russian Federation
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 461
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate Fusion XRF
CeO2, Cerium(IV) oxide (ppm) | 4376.0018 | 54.9203 | 4330.8713 | 4421.1323 | IND | IND
Fe2O3, Iron(III) oxide (wt.%) | 46.0903 | 0.7731 | 45.0622 | 47.1184 | 45.7847 | 46.3959
La2O3, Lanthanum(III) oxide (ppm) | 3190.5486 | 64.0927 | 3148.7815 | 3232.3157 | 3053.4763 | 3327.6208
Nd2O3, Neodymium(III) oxide (ppm) | 1928.2034 | 132.7159 | 1801.0235 | 2055.3834 | IND | IND
Pr6O11, Praseodymium(III,IV) oxide (ppm) | 584.6694 | 103.3662 | 483.7619 | 685.5769 | IND | IND
ThO2, Thorium dioxide (ppm) | 264.7495 | 51.8904 | 194.4351 | 335.0639 | IND | IND
Thermogravimetry
LOI, Loss On Ignition @1000°C (wt.%) | 0.99639 | 0.175 | 0.80055 | 1.1922 | 0.9397 | 1.0531
Borate / Peroxide Fusion ICP
Al2O3, Aluminium(III) oxide (wt.%) | 11.3104 | 0.39261 | 11.0327 | 11.588 | 11.1194 | 11.5013
BaO, Barium oxide (ppm) | 1037.0501 | 61.6437 | 1002.7283 | 1071.372 | 1018.3105 | 1055.7898
Be, Beryllium (ppm) | 2.3019 | 0.376 | 1.911 | 2.6927 | IND | IND
Bi, Bismuth (ppm) | 2.4852 | 0.3867 | 2.0889 | 2.8815 | IND | IND
CaO, Calcium oxide (wt.%) | 1.7609 | 0.080171 | 1.7092 | 1.8125 | 1.7212 | 1.8005
CeO2, Cerium(IV) oxide (ppm) | 4311.4355 | 162.0994 | 4221.5335 | 4401.3376 | 4193.3202 | 4429.5509
Co, Cobalt (ppm) | 11.5257 | 1.7734 | 9.6504 | 13.401 | 11.122 | 11.9294
Cr2O3, Chromium(III) oxide (ppm) | 875.4856 | 59.147 | 840.2599 | 910.7113 | 854.1616 | 896.8096
Cs, Cesium (ppm) | 0.78598 | 0.058996 | 0.74972 | 0.82224 | 0.74797 | 0.82399
Cu, Copper (ppm) | 60.1422 | 19.7586 | 40.6166 | 79.6679 | IND | IND
Dy2O3, Dysprosium(III) oxide (ppm) | 39.9771 | 1.6573 | 39.0552 | 40.899 | 38.6368 | 41.3174
Er2O3, Erbium(III) oxide (ppm) | 10.0658 | 0.33855 | 9.881 | 10.2505 | 9.6887 | 10.4428
Eu2O3, Europium(III) oxide (ppm) | 54.0907 | 2.5348 | 52.6143 | 55.5671 | 52.1777 | 56.0037
Fe2O3, Iron(III) oxide (wt.%) | 45.4263 | 1.8604 | 44.0645 | 46.7882 | 44.6595 | 46.1932
Ga, Gallium (ppm) | 50.0389 | 5.8276 | 43.2986 | 56.7792 | 47.6607 | 52.4171
Gd2O3, Gadolinium(III) oxide (ppm) | 115.7064 | 9.2992 | 109.9784 | 121.4343 | 112.3804 | 119.0324
HfO2, Hafnium dioxide (ppm) | 16.6198 | 0.71581 | 16.1609 | 17.0787 | 15.8544 | 17.3852
Ho2O3, Holmium(III) oxide (ppm) | 5.2229 | 0.35039 | 5.0153 | 5.4305 | 4.9923 | 5.4535
In, Indium (ppm) | 0.60113 | 0.07596 | 0.52363 | 0.67862 | IND | IND
K2O, Potassium oxide (wt.%) | 0.28495 | 0.039916 | 0.25603 | 0.31387 | 0.27016 | 0.29974
La2O3, Lanthanum(III) oxide (ppm) | 3154.9337 | 184.7273 | 3050.8957 | 3258.9717 | 3090.0366 | 3219.8309
Li, Lithium (ppm) | 12.0055 | 1.5327 | 10.3279 | 13.6832 | IND | IND
Lu2O3, Lutetium(III) oxide (ppm) | 0.58606 | 0.049772 | 0.55661 | 0.61551 | 0.54249 | 0.62963
MgO, Magnesium oxide (wt.%) | 1.7361 | 0.085217 | 1.6794 | 1.7929 | 1.6965 | 1.7758
MnO, Manganese oxide (ppm) | 842.6844 | 63.0822 | 795.352 | 890.0167 | 816.8559 | 868.5129
Mo, Molybdenum (ppm) | 45.6795 | 4.8335 | 41.0819 | 50.2771 | 43.3931 | 47.9659
Nb2O5, Niobium(V) oxide (ppm) | 1854.4844 | 139.6827 | 1761.5736 | 1947.3951 | 1802.669 | 1906.2997
Nd2O3, Neodymium(III) oxide (ppm) | 1899.8093 | 63.3205 | 1862.9544 | 1936.6642 | 1847.5687 | 1952.05
NiO, Nickel oxide (ppm) | 81.0513 | 6.9609 | 77.5705 | 84.532 | 67.7021 | 94.4005
P2O5, Phosphorus(V) oxide (wt.%) | 0.82204 | 0.043643 | 0.78906 | 0.85501 | 0.78734 | 0.85673
PbO, Lead oxide (ppm) | 114.4986 | 6.7968 | 107.5939 | 121.4033 | 106.965 | 122.0322
Pr6O11, Praseodymium(III,IV) oxide (ppm) | 590.314 | 25.275 | 576.5282 | 604.0998 | 571.6984 | 608.9296
Rb, Rubidium (ppm) | 13.5154 | 0.5424 | 13.2241 | 13.8068 | 13.0306 | 14.0003
S, Sulphur (ppm) | 517.6667 | 82.1688 | 432.951 | 602.3824 | IND | IND
Sb, Antimony (ppm) | 2.7613 | 0.21987 | 2.5311 | 2.9915 | 2.5308 | 2.9919
SiO2, Silicon dioxide (wt.%) | 31.4637 | 0.76968 | 30.8001 | 32.1273 | 30.9723 | 31.955
Sm2O3, Samarium(III) oxide (ppm) | 254.7024 | 8.5309 | 250.6141 | 258.7908 | 246.8114 | 262.5935
SnO2, Tin dioxide (ppm) | 32.5402 | 4.2482 | 29.7012 | 35.3792 | 30.1377 | 34.9428
SrO, Strontium oxide (ppm) | 684.6537 | 18.1191 | 674.166 | 695.1415 | 670.5914 | 698.716
Ta2O5, Tantalum(V) oxide (ppm) | 30.6885 | 1.27 | 30.0176 | 31.3594 | 29.3528 | 32.0241
Tb4O7, Terbium(III,IV) oxide (ppm) | 10.6743 | 0.3591 | 10.4918 | 10.8567 | 10.3022 | 11.0463
ThO2, Thorium dioxide (ppm) | 238.7942 | 7.751 | 233.9751 | 243.6133 | 234.2383 | 243.3502
TiO2, Titanium dioxide (wt.%) | 3.0692 | 0.091857 | 3.0072 | 3.1311 | 3.0112 | 3.1272
Tm2O3, Thulium(III) oxide (ppm) | 1.0158 | 0.097594 | 0.9618 | 1.0699 | 0.95572 | 1.076
U3O8, Uranium(V,VI) oxide (ppm) | 5.6463 | 0.22991 | 5.5264 | 5.7662 | 5.4669 | 5.8257
V2O5, Vanadium(V) oxide (ppm) | 686.6225 | 29.4125 | 670.9073 | 702.3377 | 669.6776 | 703.5674
WO3, Tungsten trioxide (ppm) | 4.8899 | 0.8429 | 4.2289 | 5.5508 | IND | IND
Y2O3, Yttrium(III) oxide (ppm) | 115.5417 | 5.5362 | 112.235 | 118.8483 | 112.2805 | 118.8029
Yb2O3, Ytterbium(III) oxide (ppm) | 5.0035 | 0.28209 | 4.8576 | 5.1494 | 4.7153 | 5.2918
ZnO, Zinc oxide (ppm) | 201.3257 | 29.0686 | 162.5883 | 240.063 | IND | IND
ZrO2, Zirconium dioxide (ppm) | 814.5465 | 36.3332 | 790.3005 | 838.7925 | 785.668 | 843.4249
4-Acid Digestion
Ag, Silver (ppm) | < 2 | IND | IND | IND | IND | IND
Al, Aluminium (wt.%) | 5.9585 | 0.18013 | 5.8591 | 6.058 | 5.8563 | 6.0607
As, Arsenic (ppm) | 33.5537 | 1.0854 | 32.9617 | 34.1457 | 32.327 | 34.7804
Ba, Barium (ppm) | 957.0984 | 44.069 | 931.1315 | 983.0654 | 935.9555 | 978.2414
Be, Beryllium (ppm) | 2.0913 | 0.24371 | 1.9592 | 2.2234 | 1.9815 | 2.2011
Bi, Bismuth (ppm) | 2.4269 | 0.15827 | 2.3321 | 2.5217 | 2.3274 | 2.5264
Ca, Calcium (wt.%) | 1.2147 | 0.090313 | 1.162 | 1.2674 | 1.1958 | 1.2336
Cd, Cadmium (ppm) | < 0.1 | IND | IND | IND | IND | IND
Ce, Cerium (ppm) | 3508.4834 | 111.9702 | 3420.3035 | 3596.6632 | 3411.4065 | 3605.5602
Co, Cobalt (ppm) | 12.1248 | 0.44431 | 11.8106 | 12.439 | 11.7568 | 12.4928
Cr, Chromium (ppm) | 523.6841 | 73.1099 | 472.7326 | 574.6355 | 507.4963 | 539.8719
Cs, Cesium (ppm) | 0.76887 | 0.066983 | 0.72791 | 0.80982 | 0.73276 | 0.80497
Cu, Copper (ppm) | 55.8138 | 2.2885 | 54.6346 | 56.993 | 54.2832 | 57.3445
Dy, Dysprosium (ppm) | 33.5055 | 0.86045 | 33.117 | 33.8941 | 32.6607 | 34.3504
Er, Erbium (ppm) | 7.4597 | 0.25253 | 7.3063 | 7.613 | 7.2377 | 7.6816
Eu, Europium (ppm) | 47.6394 | 3.2615 | 45.7237 | 49.5551 | 46.5661 | 48.7127
Fe, Iron (wt.%) | 31.5178 | 0.95936 | 30.9495 | 32.0861 | 31.0755 | 31.9601
Ga, Gallium (ppm) | 45.5202 | 8.0263 | 37.719 | 53.3215 | 44.1388 | 46.9017
Gd, Gadolinium (ppm) | 97.5977 | 3.4838 | 95.4894 | 99.706 | 95.3525 | 99.8429
Hf, Hafnium (ppm) | 7.3912 | 0.83938 | 6.7803 | 8.002 | 6.9582 | 7.8241
Ho, Holmium (ppm) | 4.0576 | 0.11474 | 3.9904 | 4.1248 | 3.9529 | 4.1623
In, Indium (ppm) | 0.56812 | 0.027058 | 0.55328 | 0.58296 | 0.5396 | 0.59664
K, Potassium (wt.%) | 0.23312 | 0.0059742 | 0.23076 | 0.23549 | 0.22291 | 0.24334
La, Lanthanum (ppm) | 2476.481 | 227.9521 | 2337.8891 | 2615.0729 | 2422.1987 | 2530.7633
Li, Lithium (ppm) | 12.0397 | 1.3787 | 11.3262 | 12.7533 | 11.6348 | 12.4447
Lu, Lutetium (ppm) | 0.38585 | 0.065651 | 0.34607 | 0.42562 | 0.36197 | 0.40972
Mg, Magnesium (wt.%) | 0.99901 | 0.10801 | 0.9352 | 1.0628 | 0.97961 | 1.0184
Mn, Manganese (ppm) | 547.2361 | 79.7765 | 495.7591 | 598.7132 | 533.448 | 561.0243
Mo, Molybdenum (ppm) | 48.2464 | 2.5361 | 46.5606 | 49.9323 | 47.2489 | 49.244
Na, Sodium (wt.%) | 0.17817 | 0.01301 | 0.17029 | 0.18606 | IND | IND
Nd, Neodymium (ppm) | 1623.483 | 79.5193 | 1579.0041 | 1667.9619 | 1566.4025 | 1680.5636
Ni, Nickel (ppm) | 71.1258 | 3.1439 | 68.9999 | 73.2516 | 68.7203 | 73.5313
P, Phosphorus (wt.%) | 0.33785 | 0.02169 | 0.324 | 0.35171 | 0.33275 | 0.34296
Pb, Lead (ppm) | 103.9184 | 5.4831 | 99.811 | 108.0258 | 101.371 | 106.4658
Pr, Praseodymium (ppm) | 482.2815 | 20.1118 | 471.2553 | 493.3077 | 469.9093 | 494.6538
Rb, Rubidium (ppm) | 13.2859 | 0.98994 | 12.7563 | 13.8155 | 12.9543 | 13.6175
Re, Rhenium (ppm) | < 0.01 | IND | IND | IND | IND | IND
S, Sulphur (ppm) | 475.4111 | 48.6238 | 446.3341 | 504.4881 | IND | IND
Sb, Antimony (ppm) | 2.4174 | 0.1736 | 2.3039 | 2.5309 | 2.2985 | 2.5362
Sc, Scandium (ppm) | 38.8492 | 2.505 | 37.3175 | 40.381 | 37.7312 | 39.9673
Sm, Samarium (ppm) | 208.2366 | 7.4192 | 204.0846 | 212.3885 | 202.7799 | 213.6932
Sn, Tin (ppm) | 21.6377 | 2.0829 | 20.2223 | 23.053 | 20.7426 | 22.5328
Sr, Strontium (ppm) | 564.1131 | 35.1659 | 542.7011 | 585.5252 | 555.5121 | 572.7142
Ta, Tantalum (ppm) | 20.9304 | 1.96 | 19.2388 | 22.6219 | 19.707 | 22.1538
Tb, Terbium (ppm) | 9.0942 | 0.50289 | 8.7602 | 9.4281 | 8.8702 | 9.3182
Te, Tellurium (ppm) | 0.34127 | 0.055579 | 0.29901 | 0.38352 | 0.28728 | 0.39525
Th, Thorium (ppm) | 206.55 | 9.605 | 201.0808 | 212.0191 | 200.0615 | 213.0384
Ti, Titanium (wt.%) | 1.0566 | 0.11708 | 0.96494 | 1.1482 | 1.0112 | 1.1019
Tl, Thallium (ppm) | 0.10384 | 0.0084602 | 0.1001 | 0.10757 | IND | IND
Tm, Thulium (ppm) | 0.70061 | 0.046188 | 0.67342 | 0.7278 | 0.67679 | 0.72443
U, Uranium (ppm) | 4.4013 | 0.16104 | 4.3111 | 4.4914 | 4.2647 | 4.5378
V, Vanadium (ppm) | 352.0608 | 28.763 | 333.494 | 370.6275 | 345.0365 | 359.085
W, Tungsten (ppm) | 2.4971 | 0.36095 | 2.2506 | 2.7436 | 2.3476 | 2.6466
Y, Yttrium (ppm) | 83.097 | 3.7487 | 81.1129 | 85.0811 | 80.5917 | 85.6022
Yb, Ytterbium (ppm) | 3.0631 | 0.14929 | 2.9764 | 3.1498 | 2.9612 | 3.165
Zn, Zinc (ppm) | 139.3186 | 13.2856 | 130.99 | 147.6472 | 136.2141 | 142.423
Zr, Zirconium (ppm) | 269.8554 | 18.0276 | 256.8571 | 282.8536 | 256.8606 | 282.8501

## Indicative Values
Table 3. Indicative Values for OREAS 461
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Al2O3 | wt.% | 11.4908 | Lu2O3 | ppm | < 20 | Ta2O5 | ppm | < 100
BaO | ppm | 1116.6667 | MgO | wt.% | 1.7596 | Tb4O7 | ppm | < 20
CaO | wt.% | 1.8154 | MnO | ppm | 904.1667 | TiO2 | wt.% | 3.0758
Cr2O3 | ppm | 628.3333 | Na2O | wt.% | 0.25017 | Tm2O3 | ppm | < 10
Dy2O3 | ppm | 46.6667 | Nb2O5 | ppm | 1863.1583 | U3O8 | ppm | < 100
Er2O3 | ppm | 20 | P2O5 | wt.% | 0.85017 | V2O5 | ppm | 728.3333
Eu2O3 | ppm | < 100 | SiO2 | wt.% | 31.6529 | WO3 | ppm | < 100
Gd2O3 | ppm | 100 | Sm2O3 | ppm | 220.48 | Y2O3 | ppm | 146.25
HfO2 | ppm | < 100 | SnO2 | ppm | < 100 | Yb2O3 | ppm | 10
Ho2O3 | ppm | 30 | SO3 | wt.% | 0.11683 | ZrO2 | ppm | 800
K2O | wt.% | 0.284 | SrO | ppm | 500
Thermogravimetry
H2O- | wt.% | 0.69667
Borate / Peroxide Fusion ICP
Ag | ppm | 15.1667 | Ge | ppm | 2.5 | Se | ppm | < 20
As | ppm | 79.3256 | Na | wt.% | 0.16624 | Te | ppm | < 1
B | ppm | 150.8333 | Re | ppm | < 0.1 | Tl | ppm | < 0.5
Cd | ppm | < 1 | Sc | ppm | 44
4-Acid Digestion
Ge | ppm | 3.5955 | Nb | ppm | 1152.7857 | Se | ppm | 6.4762

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 461
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
CeO2, ppm | 4376.0018 | 54.9203 | 4266.1612 | 4485.8424 | 4211.241 | 4540.7626 | 1.255% | 2.51% | 3.765% | 4157.2017 | 4594.8019
Fe2O3, wt.% | 46.0903 | 0.7731 | 44.5441 | 47.6365 | 43.771 | 48.4096 | 1.677% | 3.355% | 5.032% | 43.7858 | 48.3948
La2O3, ppm | 3190.5486 | 64.0927 | 3062.3631 | 3318.734 | 2998.2704 | 3382.8268 | 2.009% | 4.018% | 6.026% | 3031.0211 | 3350.076
Nd2O3, ppm | 1928.2034 | 132.7159 | 1662.7715 | 2193.6353 | 1530.0556 | 2326.3513 | 6.883% | 13.77% | 20.65% | 1831.7933 | 2024.6136
Pr6O11, ppm | 584.6694 | 103.3662 | 377.937 | 791.4018 | 274.5708 | 894.7681 | 17.68% | 35.36% | 53.04% | 555.4359 | 613.9029
ThO2, ppm | 264.7495 | 51.8904 | 160.9686 | 368.5304 | 109.0782 | 420.4208 | 19.6% | 39.2% | 58.8% | 251.512 | 277.987
Thermogravimetry
LOI, wt.% | 0.99639 | 0.175 | 0.6464 | 1.3464 | 0.4714 | 1.5214 | 17.56% | 35.13% | 52.69% | 0.94657 | 1.0462
Borate / Peroxide Fusion ICP
Al2O3, wt.% | 11.3104 | 0.39261 | 10.5251 | 12.0956 | 10.1325 | 12.4882 | 3.471% | 6.942% | 10.41% | 10.7448 | 11.8759
BaO, ppm | 1037.0501 | 61.6437 | 913.7628 | 1160.3375 | 852.1191 | 1221.9812 | 5.944% | 11.89% | 17.83% | 985.1976 | 1088.9026
Be, ppm | 2.3019 | 0.376 | 1.5499 | 3.0539 | 1.1739 | 3.4299 | 16.33% | 32.67% | 49% | 2.1868 | 2.417
Bi, ppm | 2.4852 | 0.3867 | 1.7118 | 3.2586 | 1.3251 | 3.6453 | 15.56% | 31.12% | 46.68% | 2.3609 | 2.6095
CaO, wt.% | 1.7609 | 0.080171 | 1.6005 | 1.9212 | 1.5204 | 2.0014 | 4.553% | 9.106% | 13.66% | 1.6728 | 1.8489
CeO2, ppm | 4311.4355 | 162.0994 | 3987.2368 | 4635.6343 | 3825.1375 | 4797.7336 | 3.76% | 7.52% | 11.28% | 4095.8638 | 4527.0073
Co, ppm | 11.5257 | 1.7734 | 7.9789 | 15.0724 | 6.2055 | 16.8458 | 15.39% | 30.77% | 46.16% | 10.9494 | 12.102
Cr2O3, ppm | 875.4856 | 59.147 | 757.1916 | 993.7796 | 698.0445 | 1052.9267 | 6.756% | 13.51% | 20.27% | 831.7113 | 919.2599
Cs, ppm | 0.78598 | 0.058996 | 0.66799 | 0.90397 | 0.60899 | 0.96296 | 7.506% | 15.01% | 22.52% | 0.74668 | 0.82528
Cu, ppm | 60.1422 | 19.7586 | 20.6251 | 99.6593 | 0.86656 | 119.4179 | 32.85% | 65.71% | 98.56% | 57.1351 | 63.1493
Dy2O3, ppm | 39.9771 | 1.6573 | 36.6625 | 43.2917 | 35.0052 | 44.949 | 4.146% | 8.291% | 12.44% | 37.9782 | 41.976
Er2O3, ppm | 10.0658 | 0.33855 | 9.3887 | 10.7429 | 9.0501 | 11.0814 | 3.363% | 6.727% | 10.09% | 9.5625 | 10.569
Eu2O3, ppm | 54.0907 | 2.5348 | 49.0212 | 59.1602 | 46.4864 | 61.695 | 4.686% | 9.372% | 14.06% | 51.3862 | 56.7952
Fe2O3, wt.% | 45.4263 | 1.8604 | 41.7056 | 49.1471 | 39.8452 | 51.0075 | 4.095% | 8.191% | 12.29% | 43.155 | 47.6976
Ga, ppm | 50.0389 | 5.8276 | 38.3836 | 61.6942 | 32.556 | 67.5218 | 11.65% | 23.29% | 34.94% | 47.5369 | 52.5408
Gd2O3, ppm | 115.7064 | 9.2992 | 97.108 | 134.3048 | 87.8088 | 143.604 | 8.037% | 16.07% | 24.11% | 109.9211 | 121.4917
HfO2, ppm | 16.6198 | 0.71581 | 15.1882 | 18.0515 | 14.4724 | 18.7673 | 4.307% | 8.614% | 12.92% | 15.7888 | 17.4508
Ho2O3, ppm | 5.2229 | 0.35039 | 4.5221 | 5.9237 | 4.1717 | 6.2741 | 6.709% | 13.42% | 20.13% | 4.9618 | 5.4841
In, ppm | 0.60113 | 0.07596 | 0.44921 | 0.75305 | 0.37325 | 0.82901 | 12.64% | 25.27% | 37.91% | 0.57107 | 0.63118
K2O, wt.% | 0.28495 | 0.039916 | 0.20512 | 0.36478 | 0.1652 | 0.4047 | 14.01% | 28.02% | 42.02% | 0.2707 | 0.2992
La2O3, ppm | 3154.9337 | 184.7273 | 2785.4791 | 3524.3883 | 2600.7519 | 3709.1156 | 5.855% | 11.71% | 17.57% | 2997.187 | 3312.6804
Li, ppm | 12.0055 | 1.5327 | 8.9402 | 15.0709 | 7.4075 | 16.6035 | 12.77% | 25.53% | 38.3% | 11.4052 | 12.6058
Lu2O3, ppm | 0.58606 | 0.049772 | 0.48652 | 0.6856 | 0.43674 | 0.73538 | 8.493% | 16.99% | 25.48% | 0.55676 | 0.61536
MgO, wt.% | 1.7361 | 0.085217 | 1.5657 | 1.9066 | 1.4805 | 1.9918 | 4.908% | 9.817% | 14.73% | 1.6493 | 1.8229
MnO, ppm | 842.6844 | 63.0822 | 716.5201 | 968.8487 | 653.4379 | 1031.9309 | 7.486% | 14.97% | 22.46% | 800.5502 | 884.8186
Mo, ppm | 45.6795 | 4.8335 | 36.0125 | 55.3465 | 31.179 | 60.18 | 10.58% | 21.16% | 31.74% | 43.3955 | 47.9635
Nb2O5, ppm | 1854.4844 | 139.6827 | 1575.1191 | 2133.8497 | 1435.4364 | 2273.5323 | 7.532% | 15.06% | 22.6% | 1761.7602 | 1947.2086
Nd2O3, ppm | 1899.8093 | 63.3205 | 1773.1684 | 2026.4503 | 1709.8479 | 2089.7708 | 3.333% | 6.666% | 9.999% | 1804.8189 | 1994.7998
NiO, ppm | 81.0513 | 6.9609 | 67.1294 | 94.9731 | 60.1685 | 101.934 | 8.588% | 17.18% | 25.76% | 76.9987 | 85.1038
P2O5, wt.% | 0.82204 | 0.043643 | 0.73475 | 0.90932 | 0.69111 | 0.95296 | 5.309% | 10.62% | 15.93% | 0.78093 | 0.86314
PbO, ppm | 114.4986 | 6.7968 | 100.905 | 128.0922 | 94.1082 | 134.889 | 5.936% | 11.87% | 17.81% | 108.7737 | 120.2235
Pr6O11, ppm | 590.314 | 25.275 | 539.764 | 640.864 | 514.489 | 666.1389 | 4.282% | 8.563% | 12.84% | 560.7983 | 619.8297
Rb, ppm | 13.5154 | 0.5424 | 12.4306 | 14.6002 | 11.8882 | 15.1426 | 4.013% | 8.026% | 12.04% | 12.8397 | 14.1912
S, ppm | 517.6667 | 82.1688 | 353.329 | 682.0044 | 271.1601 | 764.1732 | 15.87% | 31.75% | 47.62% | 491.7833 | 543.55
Sb, ppm | 2.7613 | 0.21987 | 2.3216 | 3.2011 | 2.1017 | 3.4209 | 7.962% | 15.92% | 23.89% | 2.6233 | 2.8994
SiO2, wt.% | 31.4637 | 0.76968 | 29.9243 | 33.003 | 29.1546 | 33.7727 | 2.446% | 4.892% | 7.339% | 29.8905 | 33.0369
Sm2O3, ppm | 254.7024 | 8.5309 | 237.6407 | 271.7642 | 229.1099 | 280.295 | 3.349% | 6.699% | 10.05% | 241.9673 | 267.4376
SnO2, ppm | 32.5402 | 4.2482 | 24.0437 | 41.0367 | 19.7955 | 45.285 | 13.06% | 26.11% | 39.17% | 30.9132 | 34.1672
SrO, ppm | 684.6537 | 18.1191 | 648.4155 | 720.8919 | 630.2965 | 739.011 | 2.646% | 5.293% | 7.939% | 650.421 | 718.8864
Ta2O5, ppm | 30.6885 | 1.27 | 28.1484 | 33.2285 | 26.8783 | 34.4986 | 4.138% | 8.277% | 12.42% | 29.154 | 32.2229
Tb4O7, ppm | 10.6743 | 0.3591 | 9.9561 | 11.3925 | 9.597 | 11.7516 | 3.364% | 6.728% | 10.09% | 10.1406 | 11.208
ThO2, ppm | 238.7942 | 7.751 | 223.2922 | 254.2962 | 215.5413 | 262.0472 | 3.246% | 6.492% | 9.738% | 226.8545 | 250.7339
TiO2, wt.% | 3.0692 | 0.091857 | 2.8854 | 3.2529 | 2.7936 | 3.3447 | 2.993% | 5.986% | 8.979% | 2.9157 | 3.2226
Tm2O3, ppm | 1.0158 | 0.097594 | 0.82066 | 1.211 | 0.72306 | 1.3086 | 9.607% | 19.21% | 28.82% | 0.96505 | 1.0666
U3O8, ppm | 5.6463 | 0.22991 | 5.1865 | 6.1061 | 4.9566 | 6.336 | 4.072% | 8.144% | 12.22% | 5.364 | 5.9286
V2O5, ppm | 686.6225 | 29.4125 | 627.7975 | 745.4475 | 598.3849 | 774.86 | 4.284% | 8.567% | 12.85% | 652.2913 | 720.9536
WO3, ppm | 4.8899 | 0.8429 | 3.2041 | 6.5757 | 2.3612 | 7.4186 | 17.24% | 34.48% | 51.71% | 4.6454 | 5.1344
Y2O3, ppm | 115.5417 | 5.5362 | 104.4692 | 126.6141 | 98.933 | 132.1503 | 4.792% | 9.583% | 14.37% | 109.7646 | 121.3188
Yb2O3, ppm | 5.0035 | 0.28209 | 4.4393 | 5.5677 | 4.1573 | 5.8498 | 5.638% | 11.28% | 16.91% | 4.7533 | 5.2537
ZnO, ppm | 201.3257 | 29.0686 | 143.1885 | 259.4628 | 114.12 | 288.5313 | 14.44% | 28.88% | 43.32% | 191.2594 | 211.3919
ZrO2, ppm | 814.5465 | 36.3332 | 741.8801 | 887.2129 | 705.5469 | 923.546 | 4.461% | 8.921% | 13.38% | 773.8192 | 855.2738
4-Acid Digestion
Ag, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Al, wt.% | 5.9585 | 0.18013 | 5.5983 | 6.3188 | 5.4181 | 6.4989 | 3.023% | 6.046% | 9.069% | 5.6606 | 6.2565
As, ppm | 33.5537 | 1.0854 | 31.3829 | 35.7245 | 30.2976 | 36.8099 | 3.235% | 6.47% | 9.704% | 31.876 | 35.2314
Ba, ppm | 957.0984 | 44.069 | 868.9604 | 1045.2365 | 824.8913 | 1089.3056 | 4.604% | 9.209% | 13.81% | 909.2435 | 1004.9534
Be, ppm | 2.0913 | 0.24371 | 1.6039 | 2.5787 | 1.3602 | 2.8224 | 11.65% | 23.31% | 34.96% | 1.9868 | 2.1959
Bi, ppm | 2.4269 | 0.15827 | 2.1104 | 2.7434 | 1.9521 | 2.9017 | 6.521% | 13.04% | 19.56% | 2.3055 | 2.5482
Ca, wt.% | 1.2147 | 0.090313 | 1.0341 | 1.3953 | 0.94376 | 1.4856 | 7.435% | 14.87% | 22.31% | 1.154 | 1.2754
Cd, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 3508.4834 | 111.9702 | 3284.543 | 3732.4237 | 3172.5728 | 3844.3939 | 3.191% | 6.383% | 9.574% | 3333.0592 | 3683.9075
Co, ppm | 12.1248 | 0.44431 | 11.2362 | 13.0134 | 10.7919 | 13.4577 | 3.665% | 7.329% | 10.99% | 11.5186 | 12.731
Cr, ppm | 523.6841 | 73.1099 | 377.4642 | 669.9039 | 304.3543 | 743.0139 | 13.96% | 27.92% | 41.88% | 497.4999 | 549.8683
Cs, ppm | 0.76887 | 0.066983 | 0.6349 | 0.90283 | 0.56792 | 0.96982 | 8.712% | 17.42% | 26.14% | 0.73042 | 0.80731
Cu, ppm | 55.8138 | 2.2885 | 51.2368 | 60.3909 | 48.9483 | 62.6794 | 4.1% | 8.201% | 12.3% | 53.0231 | 58.6045
Dy, ppm | 33.5055 | 0.86045 | 31.7847 | 35.2264 | 30.9242 | 36.0869 | 2.568% | 5.136% | 7.704% | 31.8303 | 35.1808
Er, ppm | 7.4597 | 0.25253 | 6.9546 | 7.9647 | 6.7021 | 8.2173 | 3.385% | 6.771% | 10.16% | 7.0867 | 7.8326
Eu, ppm | 47.6394 | 3.2615 | 41.1164 | 54.1624 | 37.8549 | 57.4239 | 6.846% | 13.69% | 20.54% | 45.2574 | 50.0214
Fe, wt.% | 31.5178 | 0.95936 | 29.5991 | 33.4365 | 28.6397 | 34.3959 | 3.044% | 6.088% | 9.132% | 29.9419 | 33.0937
Ga, ppm | 45.5202 | 8.0263 | 29.4677 | 61.5727 | 21.4415 | 69.599 | 17.63% | 35.26% | 52.9% | 43.2442 | 47.7963
Gd, ppm | 97.5977 | 3.4838 | 90.6301 | 104.5653 | 87.1464 | 108.049 | 3.57% | 7.139% | 10.71% | 92.7178 | 102.4776
Hf, ppm | 7.3912 | 0.83938 | 5.7124 | 9.0699 | 4.873 | 9.9093 | 11.36% | 22.71% | 34.07% | 7.0216 | 7.7607
Ho, ppm | 4.0576 | 0.11474 | 3.8281 | 4.2871 | 3.7134 | 4.4018 | 2.828% | 5.655% | 8.483% | 3.8547 | 4.2605
In, ppm | 0.56812 | 0.027058 | 0.51401 | 0.62224 | 0.48695 | 0.64929 | 4.763% | 9.525% | 14.29% | 0.53972 | 0.59653
K, wt.% | 0.23312 | 0.0059742 | 0.22118 | 0.24507 | 0.2152 | 0.25105 | 2.563% | 5.125% | 7.688% | 0.22147 | 0.24478
La, ppm | 2476.481 | 227.9521 | 2020.5769 | 2932.3852 | 1792.6248 | 3160.3373 | 9.205% | 18.41% | 27.61% | 2352.657 | 2600.3051
Li, ppm | 12.0397 | 1.3787 | 9.2824 | 14.7971 | 7.9037 | 16.1758 | 11.45% | 22.9% | 34.35% | 11.4378 | 12.6417
Lu, ppm | 0.38585 | 0.065651 | 0.25454 | 0.51715 | 0.18889 | 0.5828 | 17.01% | 34.03% | 51.04% | 0.36655 | 0.40514
Mg, wt.% | 0.99901 | 0.10801 | 0.783 | 1.215 | 0.675 | 1.323 | 10.81% | 21.62% | 32.43% | 0.94906 | 1.049
Mn, ppm | 547.2361 | 79.7765 | 387.6831 | 706.7891 | 307.9067 | 786.5656 | 14.58% | 29.16% | 43.73% | 519.8743 | 574.598
Mo, ppm | 48.2464 | 2.5361 | 43.1742 | 53.3187 | 40.6381 | 55.8548 | 5.257% | 10.51% | 15.77% | 45.8341 | 50.6588
Na, wt.% | 0.17817 | 0.01301 | 0.15215 | 0.2042 | 0.13914 | 0.21721 | 7.302% | 14.6% | 21.91% | 0.16927 | 0.18708
Nd, ppm | 1623.483 | 79.5193 | 1464.4444 | 1782.5217 | 1384.9251 | 1862.041 | 4.898% | 9.796% | 14.69% | 1542.3089 | 1704.6572
Ni, ppm | 71.1258 | 3.1439 | 64.838 | 77.4135 | 61.6942 | 80.5574 | 4.42% | 8.84% | 13.26% | 67.5695 | 74.6821
P, wt.% | 0.33785 | 0.02169 | 0.29447 | 0.38123 | 0.27278 | 0.40292 | 6.42% | 12.84% | 19.26% | 0.32096 | 0.35475
Pb, ppm | 103.9184 | 5.4831 | 92.9521 | 114.8846 | 87.469 | 120.3678 | 5.276% | 10.55% | 15.83% | 98.7225 | 109.1143
Pr, ppm | 482.2815 | 20.1118 | 442.0578 | 522.5052 | 421.946 | 542.6171 | 4.17% | 8.34% | 12.51% | 458.1675 | 506.3956
Rb, ppm | 13.2859 | 0.98994 | 11.306 | 15.2658 | 10.3161 | 16.2557 | 7.451% | 14.9% | 22.35% | 12.6216 | 13.9502
Re, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, ppm | 475.4111 | 48.6238 | 378.1635 | 572.6588 | 329.5396 | 621.2826 | 10.23% | 20.46% | 30.68% | 451.6406 | 499.1817
Sb, ppm | 2.4174 | 0.1736 | 2.0702 | 2.7646 | 1.8966 | 2.9382 | 7.181% | 14.36% | 21.54% | 2.2965 | 2.5383
Sc, ppm | 38.8492 | 2.505 | 33.8393 | 43.8591 | 31.3344 | 46.3641 | 6.448% | 12.9% | 19.34% | 36.9068 | 40.7917
Sm, ppm | 208.2366 | 7.4192 | 193.3981 | 223.075 | 185.9789 | 230.4943 | 3.563% | 7.126% | 10.69% | 197.8247 | 218.6484
Sn, ppm | 21.6377 | 2.0829 | 17.472 | 25.8034 | 15.3891 | 27.8862 | 9.626% | 19.25% | 28.88% | 20.5558 | 22.7196
Sr, ppm | 564.1131 | 35.1659 | 493.7813 | 634.445 | 458.6153 | 669.611 | 6.234% | 12.47% | 18.7% | 535.9075 | 592.3188
Ta, ppm | 20.9304 | 1.96 | 17.0103 | 24.8504 | 15.0503 | 26.8104 | 9.364% | 18.73% | 28.09% | 19.8838 | 21.9769
Tb, ppm | 9.0942 | 0.50289 | 8.0884 | 10.1 | 7.5855 | 10.6029 | 5.53% | 11.06% | 16.59% | 8.6395 | 9.5489
Te, ppm | 0.34127 | 0.055579 | 0.23011 | 0.45243 | 0.17453 | 0.50801 | 16.29% | 32.57% | 48.86% | 0.32421 | 0.35833
Th, ppm | 206.55 | 9.605 | 187.34 | 225.7599 | 177.7351 | 235.3649 | 4.65% | 9.3% | 13.95% | 196.2225 | 216.8775
Ti, wt.% | 1.0566 | 0.11708 | 0.82242 | 1.2907 | 0.70535 | 1.4078 | 11.08% | 22.16% | 33.24% | 1.0037 | 1.1094
Tl, ppm | 0.10384 | 0.0084602 | 0.086916 | 0.12076 | 0.078455 | 0.12922 | 8.148% | 16.3% | 24.44% | 0.098644 | 0.10903
Tm, ppm | 0.70061 | 0.046188 | 0.60824 | 0.79299 | 0.56205 | 0.83918 | 6.593% | 13.19% | 19.78% | 0.66558 | 0.73564
U, ppm | 4.4013 | 0.16104 | 4.0792 | 4.7233 | 3.9181 | 4.8844 | 3.659% | 7.318% | 10.98% | 4.1812 | 4.6213
V, ppm | 352.0608 | 28.763 | 294.5347 | 409.5868 | 265.7717 | 438.3498 | 8.17% | 16.34% | 24.51% | 334.4577 | 369.6638
W, ppm | 2.4971 | 0.36095 | 1.7752 | 3.219 | 1.4142 | 3.5799 | 14.45% | 28.91% | 43.36% | 2.3722 | 2.6219
Y, ppm | 83.097 | 3.7487 | 75.5996 | 90.5943 | 71.8509 | 94.343 | 4.511% | 9.022% | 13.53% | 78.9421 | 87.2518
Yb, ppm | 3.0631 | 0.14929 | 2.7645 | 3.3617 | 2.6153 | 3.511 | 4.874% | 9.748% | 14.62% | 2.91 | 3.2163
Zn, ppm | 139.3186 | 13.2856 | 112.7473 | 165.8898 | 99.4617 | 179.1754 | 9.536% | 19.07% | 28.61% | 132.3526 | 146.2845
Zr, ppm | 269.8554 | 18.0276 | 233.8002 | 305.9105 | 215.7726 | 323.9381 | 6.68% | 13.36% | 20.04% | 256.3626 | 283.3481

## Fusion XRF
| Analytical results for Al2O3 in OREAS 461 (Indicative Value 11.49 wt.%) | [8 empty] | Indicative
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 11.51 | 11.2 | 11.7 | 11.8 |  |  |  | 1
| 1 | 2 | 11.5 | 11.3 | 11.5 | 11.6 |  |  |  | 1
| 1 | 3 | 11.46 | 11.3 | 11.5 | 11.7 |  |  |  | 16
| 1 | 4 | 11.53 | 11 | 11.6 | 11.7 |  |  |  | 11.4908
| 1 | 5 | 11.55 | 11.2 | 11.4 | 11.6 |  |  |  | 7
| 1 | 6 | 11.53 | 11.4 | 11.5 | 11.7
| Mean |  | 11.5133 | 11.2333 | 11.5333 | 11.6833
| Median |  | 11.52 | 11.25 | 11.5 | 11.7
| Std Dev. |  | 0.031411 | 0.13663 | 0.10328 | 0.075277
| Rel.Std.Dev. |  | 0.2728% | 1.216% | 0.8955% | 0.6443%
| PDM3 |  | 0.1958% | -2.241% | 0.3699% | 1.675%
| Z-Score (Absolute) |  | 0.08 | 2.3 | 0.08 | 1.27
| Analytical results for BaO in OREAS 461 (Indicative Value 1117 ppm) | [8 empty] | Indicative
BaO | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 1200 | [5 empty] | 1
| 1 | 2 | N.A. | 1100 | [5 empty] | 2
| 1 | 3 | N.A. | 1100 | [5 empty] | 16
| 1 | 4 | N.A. | 1000 | [5 empty] | 1116.6667
| 1 | 5 | N.A. | 1200 | [5 empty] | 8
| 1 | 6 | -- | 1100
| Mean |  |  | 1116.6667
| Median |  |  | 1100
| Std Dev. |  |  | 75.2773
| Rel.Std.Dev. |  |  | 6.741%
| PDM3 |  |  | -2.887e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for CaO in OREAS 461 (Indicative Value 1.82 wt.%) | [8 empty] | Indicative
CaO | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.8 | 1.85 | 1.76 | 1.88 |  |  |  | 1
| 1 | 2 | 1.8 | 1.82 | 1.77 | 1.87 |  |  |  | 3
| 1 | 3 | 1.8 | 1.8 | 1.75 | 1.88 |  |  |  | 16
| 1 | 4 | 1.8 | 1.81 | 1.77 | 1.88 |  |  |  | 1.8154
| 1 | 5 | 1.79 | 1.84 | 1.74 | 1.88 |  |  |  | 9
| 1 | 6 | 1.79 | 1.84 | 1.75 | 1.9
| Mean |  | 1.7967 | 1.8267 | 1.7567 | 1.8817
| Median |  | 1.8 | 1.83 | 1.755 | 1.88
| Std Dev. |  | 0.005164 | 0.019664 | 0.012111 | 0.0098319
| Rel.Std.Dev. |  | 0.2874% | 1.076% | 0.6894% | 0.5225%
| PDM3 |  | -1.033% | 0.6197% | -3.236% | 3.649%
| Z-Score (Absolute) |  | 0.29 | 0.29 | 1.06 | 1.35
| Analytical results for CeO2 in OREAS 461 (Certified Value 4376 ppm) | [8 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 4500 | 4400 | 4400 | 4400 | 4610 | 4420 | 4300 | 1
| 1 | 2 | 4400 | 4400 | 4300 | 4300 | 4610 | 4420 | 4300 | #N/A
| 1 | 3 | 4400 | 4400 | 4090 | 4300 | 4610 | 4420 | 4400 | 16
| 1 | 4 | 4400 | 4400 | 4400 | 4400 | 4510 | 4420 | 4300 | 4376.0018
| 1 | 5 | 4400 | 4300 | 4400 | 4300 | 4610 | 4420 | 4300 | 7
| 1 | 6 | 4400 | 4400 | 4400 | 4300 | 4610 | 4420 | 4300
| Mean |  | 4416.6667 | 4383.3333 | 4331.6667 | 4333.3333 | 4593.3333 | 4420 | 4316.6667
| Median |  | 4400 | 4400 | 4400 | 4300 | 4610 | 4420 | 4300
| Std Dev. |  | 40.8248 | 40.8248 | 124.9667 | 51.6398 | 40.8248 | 0 | 40.8248
| Rel.Std.Dev. |  | 0.9243% | 0.9314% | 2.885% | 1.192% | 0.8888% | 0% | 0.9457%
| PDM3 |  | 0.9293% | 0.1675% | -1.013% | -0.9751% | 4.966% | 1.005% | -1.356%
| Z-Score (Absolute) |  | 0.58 | 0.01 | 0 | 0.88 | 3.74 | 0.67 | 1.18
| Analytical results for Cr2O3 in OREAS 461 (Indicative Value 628 ppm) | [8 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 590 | [5 empty] | 1
| 1 | 2 | N.A. | 620 | [5 empty] | 4
| 1 | 3 | N.A. | 620 | [5 empty] | 16
| 1 | 4 | N.A. | 660 | [5 empty] | 628.3333
| 1 | 5 | N.A. | 640 | [5 empty] | 10
| 1 | 6 | -- | 640
| Mean |  |  | 628.3333
| Median |  |  | 630
| Std Dev. |  |  | 24.0139
| Rel.Std.Dev. |  |  | 3.822%
| PDM3 |  |  | 4.441e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Dy2O3 in OREAS 461 (Indicative Value 46.7 ppm) | [8 empty] | Indicative
Dy2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | 195 | N.A. | 50 |  |  |  |  | 1
| 1 | 2 | 195 | N.A. | 40 |  |  |  |  | 5
| 1 | 3 | 195 | N.A. | 50 |  |  |  |  | 16
| 1 | 4 | 195 | N.A. | 50 |  |  |  |  | 46.6667
| 1 | 5 | 195 | N.A. | 40 |  |  |  |  | 11
| 1 | 6 | 195 | -- | 50
| Mean |  | 195 |  | 46.6667
| Median |  | 195 |  | 50
| Std Dev. |  | 0 |  | 5.164
| Rel.Std.Dev. |  | 0% |  | 11.07%
| PDM3 |  | 317.9% |  | -7.772e-14%
| Z-Score (Absolute) |  | 0.67 | NA | 0.67
| Analytical results for Er2O3 in OREAS 461 (Indicative Value 20 ppm) | [8 empty] | Indicative
Er2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 100 | N.A. | 20 |  |  |  |  | 1
| 1 | 2 | < 100 | N.A. | < 10 |  |  |  |  | 6
| 1 | 3 | < 100 | N.A. | 20 |  |  |  |  | 16
| 1 | 4 | < 100 | N.A. | 20 |  |  |  |  | 20
| 1 | 5 | < 100 | N.A. | 20 |  |  |  |  | 12
| 1 | 6 | < 100 | -- | 20
| Mean |  |  |  | 20
| Median |  |  |  | 20
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  |  | 0%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Eu2O3 in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Eu2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 09 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 100 | < 100 | < 100 |  |  |  |  | 1
| 1 | 2 | < 100 | < 100 | < 100 |  |  |  |  | 7
| 1 | 3 | < 100 | < 100 | < 100 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | < 100 |  |  |  |  | < 100
| 1 | 5 | < 100 | < 100 | < 100 |  |  |  |  | 13
| 1 | 6 | < 100 | < 100 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Fe2O3 in OREAS 461 (Certified Value 46.09 wt.%) | [8 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 09 | 13 | 14 | 15 | 22 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 46.1 | 47.4 | 45.3 | 45.3 | 46.4 |  |  | 1
| 1 | 2 | 46.37 | 47 | 45.2 | 45.2 | 46.5 |  |  | #N/A
| 1 | 3 | 46.51 | 46.8 | 45.4 | 45.2 | 46.4 |  |  | 16
| 1 | 4 | 46.5 | 46.9 | 45.3 | 45.1 | 46.3 |  |  | 46.0903
| 1 | 5 | 46.32 | 47.3 | 45.1 | 45.2 | 46.3 |  |  | 8
| 1 | 6 | 46.41 | 47.7 | 45.3 | 45.5 | 46.4
| Mean |  | 46.3683 | 47.1833 | 45.2667 | 45.25 | 46.3833
| Median |  | 46.39 | 47.15 | 45.3 | 45.2 | 46.4
| Std Dev. |  | 0.15065 | 0.34303 | 0.10328 | 0.13784 | 0.075277
| Rel.Std.Dev. |  | 0.3249% | 0.727% | 0.2282% | 0.3046% | 0.1623%
| PDM3 |  | 0.6032% | 2.371% | -1.787% | -1.823% | 0.6357%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.91 | 0.93 | 0.01
| Analytical results for Gd2O3 in OREAS 461 (Indicative Value 100 ppm) | [8 empty] | Indicative
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 09 | 14 | 22 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 100 | 100 | < 100 | 100 |  |  |  | 1
| 1 | 2 | 100 | 200 | < 100 | 100 |  |  |  | 8
| 1 | 3 | 100 | 100 | < 100 | 100 |  |  |  | 16
| 1 | 4 | 100 | 100 | 100 | 100 |  |  |  | 100
| 1 | 5 | 100 | 200 | < 100 | 100 |  |  |  | 14
| 1 | 6 | 100 | 100 | 100 | 100
| Mean |  | 100 | 133.3333 | 100 | 100
| Median |  | 100 | 100 | 100 | 100
| Std Dev. |  | 0 | 51.6398 | 0 | 0
| Rel.Std.Dev. |  | 0% | 38.73% | 0% | 0%
| PDM3 |  | 0% | 33.33% | 0% | 0%
| Z-Score (Absolute) |  | 0 | 1.35 | 1.35 | 0
| Analytical results for HfO2 in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
HfO2 | Round | Replicate | Lab | [6 empty] | 1
| No | No | 13 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | [6 empty] | 1
| 1 | 2 | < 100 | [6 empty] | 9
| 1 | 3 | < 100 | [6 empty] | 16
| 1 | 4 | < 100 | [6 empty] | < 100
| 1 | 5 | < 100 | [6 empty] | 15
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho2O3 in OREAS 461 (Indicative Value 30 ppm) | [8 empty] | Indicative
Ho2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | 100 | 10 | [5 empty] | 1
| 1 | 2 | < 100 | < 10 | [5 empty] | 10
| 1 | 3 | < 100 | 10 | [5 empty] | 16
| 1 | 4 | < 100 | 10 | [5 empty] | 30
| 1 | 5 | < 100 | 10 | [5 empty] | 16
| 1 | 6 | < 100 | 10
| Mean |  | 100 | 10
| Median |  | 100 | 10
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  | 233.3% | -66.67%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for K2O in OREAS 461 (Indicative Value 0.284 wt.%) | [8 empty] | Indicative
K2O | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.28 | 0.4 | 0.28 | 0.29 |  |  |  | 1
| 1 | 2 | 0.28 | 0.38 | 0.28 | 0.29 |  |  |  | 11
| 1 | 3 | 0.28 | 0.36 | 0.28 | 0.29 |  |  |  | 16
| 1 | 4 | 0.28 | 0.38 | 0.29 | 0.29 |  |  |  | 0.284
| 1 | 5 | 0.28 | 0.38 | 0.28 | 0.29 |  |  |  | 17
| 1 | 6 | 0.28 | 0.39 | 0.27 | 0.29
| Mean |  | 0.28 | 0.38167 | 0.28 | 0.29
| Median |  | 0.28 | 0.38 | 0.28 | 0.29
| Std Dev. |  | 0 | 0.013292 | 0.0063246 | 0
| Rel.Std.Dev. |  | 0% | 3.483% | 2.259% | 0%
| PDM3 |  | -1.408% | 34.39% | -1.408% | 2.113%
| Z-Score (Absolute) |  | 0.81 | 12.9 | 0.54 | 0.54
| Analytical results for La2O3 in OREAS 461 (Certified Value 3191 ppm) | [8 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 3100 | 3200 | 3300 | 3200 | 3200 | 3280 | 3200 | 1
| 1 | 2 | 3100 | 3200 | 3200 | 3200 | 3200 | 3280 | 3100 | #N/A
| 1 | 3 | 3200 | 3200 | 3200 | 3100 | 3100 | 3280 | 3200 | 16
| 1 | 4 | 3100 | 3200 | 3100 | 3000 | 3200 | 3280 | 3200 | 3190.5486
| 1 | 5 | 3200 | 3200 | 3200 | 3200 | 3200 | 3280 | 3200 | 9
| 1 | 6 | 3200 | 3200 | 3100 | 3200 | 3200 | 3280 | 3200
| Mean |  | 3150 | 3200 | 3183.3333 | 3150 | 3183.3333 | 3280 | 3183.3333
| Median |  | 3150 | 3200 | 3200 | 3200 | 3200 | 3280 | 3200
| Std Dev. |  | 54.7723 | 0 | 75.2773 | 83.666 | 40.8248 | 0 | 40.8248
| Rel.Std.Dev. |  | 1.739% | 0% | 2.365% | 2.656% | 1.282% | 0% | 1.282%
| PDM3 |  | -1.271% | 0.2962% | -0.2261% | -1.271% | -0.2261% | 2.804% | -0.2261%
| Z-Score (Absolute) |  | 1.35 | 0.67 | 0 | 1.35 | 0 | 4.07 | 0
| Analytical results for Lu2O3 in OREAS 461 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Lu2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | < 20 | [5 empty] | 1
| 1 | 2 | < 100 | < 20 | [5 empty] | 1
| 1 | 3 | < 100 | < 20 | [5 empty] | 16
| 1 | 4 | < 100 | < 20 | [5 empty] | < 20
| 1 | 5 | < 100 | < 20 | [5 empty] | 7
| 1 | 6 | < 100 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for MgO in OREAS 461 (Indicative Value 1.76 wt.%) | [8 empty] | Indicative
MgO | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.79 | 1.69 | 1.78 | 1.77 |  |  |  | 1
| 1 | 2 | 1.8 | 1.72 | 1.79 | 1.77 |  |  |  | 2
| 1 | 3 | 1.82 | 1.68 | 1.77 | 1.79 |  |  |  | 16
| 1 | 4 | 1.81 | 1.71 | 1.76 | 1.73 |  |  |  | 1.7596
| 1 | 5 | 1.79 | 1.68 | 1.78 | 1.76 |  |  |  | 8
| 1 | 6 | 1.8 | 1.7 | 1.74 | 1.8
| Mean |  | 1.8017 | 1.6967 | 1.77 | 1.77
| Median |  | 1.8 | 1.695 | 1.775 | 1.77
| Std Dev. |  | 0.01169 | 0.01633 | 0.017889 | 0.024495
| Rel.Std.Dev. |  | 0.6489% | 0.9625% | 1.011% | 1.384%
| PDM3 |  | 2.392% | -3.576% | 0.592% | 0.592%
| Z-Score (Absolute) |  | 1.35 | 3.12 | 0 | 0
| Analytical results for MnO in OREAS 461 (Indicative Value 904 ppm) | [8 empty] | Indicative
MnO | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 900 | 1000 | 800 | 1000 |  |  |  | 1
| 1 | 2 | 900 | 900 | 800 | 1000 |  |  |  | 3
| 1 | 3 | 900 | 1000 | 800 | 900 |  |  |  | 16
| 1 | 4 | 900 | 900 | 800 | 900 |  |  |  | 904.1667
| 1 | 5 | 900 | 900 | 800 | 1000 |  |  |  | 9
| 1 | 6 | 900 | 1000 | 800 | 1000
| Mean |  | 900 | 950 | 800 | 966.6667
| Median |  | 900 | 950 | 800 | 1000
| Std Dev. |  | 0 | 54.7723 | 0 | 51.6398
| Rel.Std.Dev. |  | 0% | 5.766% | 0% | 5.342%
| PDM3 |  | -0.4608% | 5.069% | -11.52% | 6.912%
| Z-Score (Absolute) |  | 0.51 | 0.51 | 2.53 | 0.84
| Analytical results for Na2O in OREAS 461 (Indicative Value 0.25 wt.%) | [8 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.26 | 0.25 | 0.24 | 0.27 |  |  |  | 1
| 1 | 2 | 0.23 | 0.24 | 0.31 | 0.24 |  |  |  | 4
| 1 | 3 | 0.26 | 0.26 | 0.28 | 0.24 |  |  |  | 16
| 1 | 4 | 0.23 | 0.27 | 0.25 | 0.25 |  |  |  | 0.25017
| 1 | 5 | 0.24 | 0.23 | 0.24 | 0.25 |  |  |  | 10
| 1 | 6 | 0.26 | 0.24 | 0.26 | 0.26
| Mean |  | 0.24667 | 0.24833 | 0.26333 | 0.25167
| Median |  | 0.25 | 0.245 | 0.255 | 0.25
| Std Dev. |  | 0.015055 | 0.01472 | 0.027325 | 0.01169
| Rel.Std.Dev. |  | 6.104% | 5.927% | 10.38% | 4.645%
| PDM3 |  | -1.399% | -0.7328% | 5.263% | 0.5996%
| Z-Score (Absolute) |  | 0.9 | 0.45 | 1.08 | 0.45
| Analytical results for Nb2O5 in OREAS 461 (Indicative Value 1863 ppm) | [8 empty] | Indicative
Nb2O5 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | N.A. | 1800 | 1860 |  |  |  |  | 1
| 1 | 2 | N.A. | 1900 | 1860 |  |  |  |  | 5
| 1 | 3 | N.A. | 1800 | 1860 |  |  |  |  | 16
| 1 | 4 | N.A. | 1900 | 1860 |  |  |  |  | 1863.1583
| 1 | 5 | N.A. | 1900 | 1860 |  |  |  |  | 11
| 1 | 6 | -- | 1900 | 1860
| Mean |  |  | 1866.6667 | 1860
| Median |  |  | 1900 | 1860
| Std Dev. |  |  | 51.6398 | 0
| Rel.Std.Dev. |  |  | 2.766% | 0%
| PDM3 |  |  | 0.1883% | -0.1695%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Nd2O3 in OREAS 461 (Certified Value 1928 ppm) | [8 empty] | Certified
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 1995 | 1800 | 1600 | 2000 | 2100 | 1980 | 1870 | 1
| 1 | 2 | 1995 | 1900 | 1900 | 1800 | 2200 | 1980 | 1880 | #N/A
| 1 | 3 | 1995 | 1900 | 1700 | 2000 | 2200 | 1980 | 1900 | 16
| 1 | 4 | 1995 | 1900 | 1700 | 2000 | 2100 | 1980 | 1900 | 1928.2034
| 1 | 5 | 1995 | 1900 | 1700 | 1800 | 2100 | 1980 | 1890 | 10
| 1 | 6 | 1995 | 1900 | 1700 | 1900 | 2100 | 1980 | 1900
| Mean |  | 1995 | 1883.3333 | 1716.6667 | 1916.6667 | 2133.3333 | 1980 | 1890
| Median |  | 1995 | 1900 | 1700 | 1950 | 2100 | 1980 | 1895
| Std Dev. |  | 0 | 40.8248 | 98.3192 | 98.3192 | 51.6398 | 0 | 12.6491
| Rel.Std.Dev. |  | 0% | 2.168% | 5.727% | 5.13% | 2.421% | 0% | 0.6693%
| PDM3 |  | 3.464% | -2.327% | -10.97% | -0.5983% | 10.64% | 2.686% | -1.981%
| Z-Score (Absolute) |  | 0.79 | 0.17 | 2.41 | 0 | 2.21 | 0.67 | 0.27
| Analytical results for P2O5 in OREAS 461 (Indicative Value 0.85 wt.%) | [8 empty] | Indicative
P2O5 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 09 | 14 | 15 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 0.84 | 0.846 | 0.86 |  |  |  |  | 1
| 1 | 2 | 0.838 | 0.84 | 0.86 |  |  |  |  | 6
| 1 | 3 | 0.836 | 0.853 | 0.86 |  |  |  |  | 16
| 1 | 4 | 0.838 | 0.848 | 0.86 |  |  |  |  | 0.85017
| 1 | 5 | 0.843 | 0.853 | 0.87 |  |  |  |  | 12
| 1 | 6 | 0.835 | 0.853 | 0.87
| Mean |  | 0.83833 | 0.84883 | 0.86333
| Median |  | 0.838 | 0.8505 | 0.86
| Std Dev. |  | 0.0028752 | 0.0052694 | 0.005164
| Rel.Std.Dev. |  | 0.343% | 0.6208% | 0.5981%
| PDM3 |  | -1.392% | -0.1568% | 1.549%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.93
| Analytical results for Pr6O11 in OREAS 461 (Certified Value 585 ppm) | [8 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 616 | 600 | 520 | 400 | 620 | 480 | 700 | 1
| 1 | 2 | 616 | 600 | 1140 | 400 | 620 | 600 | 600 | #N/A
| 1 | 3 | 616 | 600 | 1030 | 400 | 720 | 600 | 700 | 16
| 1 | 4 | 616 | 500 | 1240 | 400 | 620 | 600 | 700 | 584.6694
| 1 | 5 | 520 | 600 | 520 | 400 | 520 | 600 | 600 | 11
| 1 | 6 | 616 | 600 | 830 | 400 | 830 | 600 | 700
| Mean |  | 600 | 583.3333 | 880 | 400 | 655 | 580 | 666.6667
| Median |  | 616 | 600 | 930 | 400 | 620 | 600 | 700
| Std Dev. |  | 39.1918 | 40.8248 | 310.2902 | 0 | 106.5364 | 48.9898 | 51.6398
| Rel.Std.Dev. |  | 6.532% | 6.999% | 35.26% | 0% | 16.27% | 8.447% | 7.746%
| PDM3 |  | 2.622% | -0.2285% | 50.51% | -31.59% | 12.03% | -0.7986% | 14.02%
| Z-Score (Absolute) |  | 0.05 | 0.28 | 3.71 | 2.77 | 0.67 | 0 | 0.85
| Analytical results for SiO2 in OREAS 461 (Indicative Value 31.65 wt.%) | [8 empty] | Indicative
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 31.66 | 32.2 | 31.6 | 31.6 |  |  |  | 1
| 1 | 2 | 31.67 | 31.9 | 31.7 | 31.5 |  |  |  | 7
| 1 | 3 | 31.68 | 31.5 | 31.7 | 31.6 |  |  |  | 16
| 1 | 4 | 31.72 | 31 | 32.2 | 31.4 |  |  |  | 31.6529
| 1 | 5 | 31.53 | 31.7 | 31.5 | 31.5 |  |  |  | 13
| 1 | 6 | 31.51 | 32 | 31.8 | 31.5
| Mean |  | 31.6283 | 31.7167 | 31.75 | 31.5167
| Median |  | 31.665 | 31.8 | 31.7 | 31.5
| Std Dev. |  | 0.086583 | 0.42622 | 0.2429 | 0.075277
| Rel.Std.Dev. |  | 0.2738% | 1.344% | 0.765% | 0.2388%
| PDM3 |  | -0.07767% | 0.2014% | 0.3067% | -0.4305%
| Z-Score (Absolute) |  | 0.49 | 0.49 | 0.86 | 1.73
| Analytical results for Sm2O3 in OREAS 461 (Indicative Value 220 ppm) | [8 empty] | Indicative
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 14 | 15 | 16 | 22 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 104 | 300 | 200 | < 500 | 230 | 200 |  | 1
| 1 | 2 | < 104 | 300 | 200 | < 500 | 230 | 200 |  | 8
| 1 | 3 | 104 | 300 | 200 | < 500 | 230 | 200 |  | 16
| 1 | 4 | 104 | 200 | 200 | < 500 | 230 | 200 |  | 220.48
| 1 | 5 | 104 | 200 | 100 | < 500 | 230 | 200 |  | 14
| 1 | 6 | 104 | 200 | 200 | < 500 | 230 | 200
| Mean |  | 104 | 250 | 183.3333 |  | 230 | 200
| Median |  | 104 | 250 | 200 |  | 230 | 200
| Std Dev. |  | 0 | 54.7723 | 40.8248 |  | 0 | 0
| Rel.Std.Dev. |  | 0% | 21.91% | 22.27% |  | 0% | 0%
| PDM3 |  | -52.83% | 13.39% | -16.85% |  | 4.318% | -9.289%
| Z-Score (Absolute) |  | 2.39 | 1.06 | 0 | NA | 0.67 | 0
| Analytical results for SnO2 in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
SnO2 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | < 100 | [5 empty] | 1
| 1 | 2 | N.A. | < 100 | [5 empty] | 9
| 1 | 3 | N.A. | < 100 | [5 empty] | 16
| 1 | 4 | N.A. | < 100 | [5 empty] | < 100
| 1 | 5 | N.A. | < 100 | [5 empty] | 15
| 1 | 6 | -- | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for SO3 in OREAS 461 (Indicative Value 0.117 wt.%) | [8 empty] | Indicative
SO3 | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.121 | [6 empty] | 1
| 1 | 2 | 0.111 | [6 empty] | 10
| 1 | 3 | 0.115 | [6 empty] | 16
| 1 | 4 | 0.11 | [6 empty] | 0.11683
| 1 | 5 | 0.118 | [6 empty] | 16
| 1 | 6 | 0.126
| Mean |  | 0.11683
| Median |  | 0.1165
| Std Dev. |  | 0.0061128
| Rel.Std.Dev. |  | 5.232%
| PDM3 |  | 2.887e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for SrO in OREAS 461 (Indicative Value 500 ppm) | [8 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 500 | [5 empty] | 1
| 1 | 2 | N.A. | 500 | [5 empty] | 11
| 1 | 3 | N.A. | 600 | [5 empty] | 16
| 1 | 4 | N.A. | 400 | [5 empty] | 500
| 1 | 5 | N.A. | 500 | [5 empty] | 17
| 1 | 6 | -- | 500
| Mean |  |  | 500
| Median |  |  | 500
| Std Dev. |  |  | 63.2456
| Rel.Std.Dev. |  |  | 12.65%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ta2O5 in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Ta2O5 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | < 100 | [5 empty] | 1
| 1 | 2 | N.A. | < 100 | [5 empty] | 1
| 1 | 3 | N.A. | < 100 | [5 empty] | 16
| 1 | 4 | N.A. | < 100 | [5 empty] | < 100
| 1 | 5 | N.A. | < 100 | [5 empty] | 7
| 1 | 6 | -- | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Tb4O7 in OREAS 461 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Tb4O7 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | < 20 | [5 empty] | 1
| 1 | 2 | < 100 | < 20 | [5 empty] | 2
| 1 | 3 | < 100 | < 20 | [5 empty] | 16
| 1 | 4 | < 100 | < 20 | [5 empty] | < 20
| 1 | 5 | < 100 | < 20 | [5 empty] | 8
| 1 | 6 | < 100 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for ThO2 in OREAS 461 (Certified Value 265 ppm) | [8 empty] | Certified
ThO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 205 | 300 | < 300 | 300 | 320 | 193 |  | 1
| 1 | 2 | 216 | 300 | < 300 | 300 | 330 | 193 |  | #N/A
| 1 | 3 | 228 | 300 | 400 | 300 | 360 | 182 |  | 16
| 1 | 4 | 205 | 300 | < 300 | 300 | 320 | 193 |  | 264.7495
| 1 | 5 | 228 | 300 | < 300 | 300 | 310 | 193 |  | 12
| 1 | 6 | 216 | 300 | < 300 | 300 | 300 | 193
| Mean |  | 216.3333 | 300 | 400 | 300 | 323.3333 | 191.1667
| Median |  | 216 | 300 | 400 | 300 | 320 | 193
| Std Dev. |  | 10.2892 | 0 |  | 0 | 20.6559 | 4.4907
| Rel.Std.Dev. |  | 4.756% | 0% |  | 0% | 6.388% | 2.349%
| PDM3 |  | -18.29% | 13.31% | 51.09% | 13.31% | 22.13% | -27.79%
| Z-Score (Absolute) |  | 0.57 | 0.57 | 1.46 | 0.57 | 0.78 | 0.9
| Analytical results for TiO2 in OREAS 461 (Indicative Value 3.08 wt.%) | [8 empty] | Indicative
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 3.08 | 3.09 | 3.01 | 3.17 |  |  |  | 1
| 1 | 2 | 3.06 | 3.06 | 3.03 | 3.16 |  |  |  | 3
| 1 | 3 | 3.07 | 3.03 | 3.02 | 3.17 |  |  |  | 16
| 1 | 4 | 3.07 | 3 | 3.02 | 3.16 |  |  |  | 3.0758
| 1 | 5 | 3.06 | 3.06 | 2.99 | 3.16 |  |  |  | 9
| 1 | 6 | 3.06 | 3.07 | 3.05 | 3.17
| Mean |  | 3.0667 | 3.0517 | 3.02 | 3.165
| Median |  | 3.065 | 3.06 | 3.02 | 3.165
| Std Dev. |  | 0.008165 | 0.031885 | 0.02 | 0.0054772
| Rel.Std.Dev. |  | 0.2662% | 1.045% | 0.6623% | 0.1731%
| PDM3 |  | -0.298% | -0.7857% | -1.815% | 2.899%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 1.13 | 3.06
| Analytical results for Tm2O3 in OREAS 461 (Indicative Value < 10 ppm) | [8 empty] | Indicative
Tm2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [5 empty] | 1
| 1 | 2 | < 100 | < 10 | [5 empty] | 4
| 1 | 3 | < 100 | < 10 | [5 empty] | 16
| 1 | 4 | < 100 | < 10 | [5 empty] | < 10
| 1 | 5 | < 100 | < 10 | [5 empty] | 10
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U3O8 in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
U3O8 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 300 | N.A. | 100 |  |  |  | 1
| 1 | 2 | < 100 | < 300 | N.A. | < 100 |  |  |  | 5
| 1 | 3 | < 100 | < 300 | N.A. | < 100 |  |  |  | 16
| 1 | 4 | < 100 | < 300 | N.A. | < 100 |  |  |  | < 100
| 1 | 5 | < 100 | < 300 | N.A. | < 100 |  |  |  | 11
| 1 | 6 | < 100 | < 300 | -- | < 100
| Mean |  |  |  |  | 100
| Median |  |  |  |  | 100
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for V2O5 in OREAS 461 (Indicative Value 728 ppm) | [8 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 740 | [5 empty] | 1
| 1 | 2 | N.A. | 710 | [5 empty] | 6
| 1 | 3 | N.A. | 740 | [5 empty] | 16
| 1 | 4 | N.A. | 720 | [5 empty] | 728.3333
| 1 | 5 | N.A. | 740 | [5 empty] | 12
| 1 | 6 | -- | 720
| Mean |  |  | 728.3333
| Median |  |  | 730
| Std Dev. |  |  | 13.2916
| Rel.Std.Dev. |  |  | 1.825%
| PDM3 |  |  | 4.441e-14%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for WO3 in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
WO3 | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | [6 empty] | 1
| 1 | 2 | < 100 | [6 empty] | 7
| 1 | 3 | < 100 | [6 empty] | 16
| 1 | 4 | < 100 | [6 empty] | < 100
| 1 | 5 | < 100 | [6 empty] | 13
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y2O3 in OREAS 461 (Indicative Value 146 ppm) | [8 empty] | Indicative
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | < 130 | < 100 | < 100 | 200 | 200 | < 130 | 120 | 1
| 1 | 2 | < 130 | 100 | < 100 | 200 | 100 | < 130 | 120 | 8
| 1 | 3 | < 130 | < 100 | < 100 | 200 | 100 | < 130 | 120 | 16
| 1 | 4 | < 130 | 100 | < 100 | 200 | < 100 | < 130 | 90 | 146.25
| 1 | 5 | < 130 | 200 | < 100 | 100 | 200 | < 130 | 90 | 14
| 1 | 6 | < 130 | 200 | < 100 | 200 | 300 | < 130 | 120
| Mean |  |  | 150 |  | 183.3333 | 180 |  | 110
| Median |  |  | 150 |  | 200 | 200 |  | 120
| Std Dev. |  |  | 57.735 |  | 40.8248 | 83.666 |  | 15.4919
| Rel.Std.Dev. |  |  | 38.49% |  | 22.27% | 46.48% |  | 14.08%
| PDM3 |  |  | 2.564% |  | 25.36% | 23.08% |  | -24.79%
| Z-Score (Absolute) |  | NA | 0.58 | NA | 1.74 | 0.58 | NA | 0.77
| Analytical results for Yb2O3 in OREAS 461 (Indicative Value 10 ppm) | [8 empty] | Indicative
Yb2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 100 | N.A. | 10 |  |  |  |  | 1
| 1 | 2 | < 100 | N.A. | < 10 |  |  |  |  | 9
| 1 | 3 | < 100 | N.A. | 10 |  |  |  |  | 16
| 1 | 4 | < 100 | N.A. | 10 |  |  |  |  | 10
| 1 | 5 | < 100 | N.A. | 10 |  |  |  |  | 15
| 1 | 6 | < 100 | -- | 10
| Mean |  |  |  | 10
| Median |  |  |  | 10
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  |  | 0%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for ZrO2 in OREAS 461 (Indicative Value 800 ppm) | [8 empty] | Indicative
ZrO2 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 800 | [5 empty] | 1
| 1 | 2 | N.A. | 800 | [5 empty] | 10
| 1 | 3 | N.A. | 800 | [5 empty] | 16
| 1 | 4 | N.A. | 800 | [5 empty] | 800
| 1 | 5 | N.A. | 800 | [5 empty] | 16
| 1 | 6 | -- | 800
| Mean |  |  | 800
| Median |  |  | 800
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 461 (Indicative Value 0.697 wt.%) | [8 empty] | Indicative
H2O- | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | at 1000°C | [6 empty] | 3
3
| 1 | 1 | 0.68 | [6 empty] | 1
| 1 | 2 | 0.73 | [6 empty] | 13
| 1 | 3 | 0.71 | [6 empty] | 16
| 1 | 4 | 0.7 | [6 empty] | 0.69667
| 1 | 5 | 0.69 | [6 empty] | 19
| 1 | 6 | 0.67
| Mean |  | 0.69667
| Median |  | 0.695
| Std Dev. |  | 0.021602
| Rel.Std.Dev. |  | 3.101%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI in OREAS 461 (Certified Value 0.996 wt.%) | [8 empty] | Certified
LOI | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 18 | wt.%
|  |  | at 1000°C | at 1000°C | at 1000°C | BF*OES | at 1000°C | BF*XRF | at 1000°C | 3
3
| 1 | 1 | 0.92 | 0.84 | 0.99 | 1.12 | 1.3 | 0.776 | 1.74 | 1
| 1 | 2 | 0.93 | 0.8 | 1.01 | 1.13 | 1.3 | 0.832 | 1.7793 | #N/A
| 1 | 3 | 0.94 | 0.83 | 0.99 | 1.13 | 1.28 | 0.83 | 1.7379 | 16
| 1 | 4 | 0.94 | 0.9 | 0.96 | 1.01 | 1.31 | 0.823 | 1.6371 | 0.99639
| 1 | 5 | 0.98 | 0.82 | 0.96 | 1.2 | 1.28 | 0.812 | 1.679 | 14
| 1 | 6 | 0.94 | 0.83 | 0.97 | 1.12 | 1.25 | 0.763 | 1.6793
| Mean |  | 0.94167 | 0.83667 | 0.98 | 1.1183 | 1.2867 | 0.806 | 1.7088
| Median |  | 0.94 | 0.83 | 0.98 | 1.125 | 1.29 | 0.8175 | 1.7086
| Std Dev. |  | 0.020412 | 0.033862 | 0.02 | 0.061128 | 0.021602 | 0.029414 | 0.052334
| Rel.Std.Dev. |  | 2.168% | 4.047% | 2.041% | 5.466% | 1.679% | 3.649% | 3.063%
| PDM3 |  | -5.492% | -16.03% | -1.645% | 12.24% | 29.13% | -19.11% | 71.5%
| Z-Score (Absolute) |  | 0.16 | 0.66 | 0 | 0.67 | 1.29 | 0.73 | 3.07
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 461 (Indicative Value 15.2 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 20 | 22 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 5 | 18 | < 5 | [12 empty] | 1
| 1 | 2 | < 5 | 14 | < 5 | [12 empty] | 15
| 1 | 3 | < 5 | 14 | < 5 | [12 empty] | 16
| 1 | 4 | < 5 | 14 | < 5 | [12 empty] | 15.1667
| 1 | 5 | < 5 | 16 | < 5 | [12 empty] | 21
| 1 | 6 | < 5 | 15 | < 5
| Mean |  |  | 15.1667
| Median |  |  | 14.5
| Std Dev. |  |  | 1.6021
| Rel.Std.Dev. |  |  | 10.56%
| PDM3 |  |  | -2.22e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Al2O3 in OREAS 461 (Certified Value 11.31 wt.%) | [16 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 11.4201 | 11.734 | 11.261 | 11.3092 | 11.34 | 10.7428 | 10.789 | 11.431 | 11.772 | 11.62 | [5 empty] | 1
| 1 | 2 | 11.4523 | 11.658 | 11.299 | 11.4441 | 11.64 | 10.6797 | 10.543 | 11.11 | 11.866 | 11.602 | [5 empty] | 1
| 1 | 3 | 11.4182 | 11.545 | 11.243 | 11.5359 | 11.41 | 10.7642 | 10.506 | 10.808 | 11.602 | 11.016 | [5 empty] | 16
| 1 | 4 | 11.3332 | 11.696 | 11.299 | 11.1917 | 11.63 | 10.8177 | 10.657 | 10.959 | 11.734 | 11.715 | [5 empty] | 11.3104
| 1 | 5 | 11.3502 | 11.602 | 11.394 | 11.4885 | 11.43 | 10.7214 | 10.411 | 11.148 | 11.961 | 11.583 | [5 empty] | 16
| 1 | 6 | 11.2671 | 11.356 | 11.261 | 11.3589 | 11.33 | 10.8819 | 10.6 | 11.45 | 12.112 | 11.696
| Mean |  | 11.3735 | 11.5985 | 11.2928 | 11.388 | 11.4633 | 10.7679 | 10.5843 | 11.151 | 11.8412 | 11.5387
| Median |  | 11.3842 | 11.63 | 11.28 | 11.4015 | 11.42 | 10.7535 | 10.5715 | 11.129 | 11.819 | 11.611
| Std Dev. |  | 0.06909 | 0.13645 | 0.054444 | 0.12706 | 0.13852 | 0.072203 | 0.13062 | 0.25456 | 0.17989 | 0.26139
| Rel.Std.Dev. |  | 0.6075% | 1.176% | 0.4821% | 1.116% | 1.208% | 0.6705% | 1.234% | 2.283% | 1.519% | 2.265%
| PDM3 |  | 0.5583% | 2.548% | -0.155% | 0.6868% | 1.352% | -4.796% | -6.419% | -1.409% | 4.693% | 2.019%
| Z-Score (Absolute) |  | 0.02 | 0.66 | 0.27 | 0.02 | 0.25 | 1.85 | 2.4 | 0.69 | 1.39 | 0.79
| Analytical results for As in OREAS 461 (Indicative Value 79 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 37 | < 100 | 79.8798 | 100 | 40 | 140 | [9 empty] | 1
| 1 | 2 | 39 | < 100 | 80.4216 | 100 | 50 | 150 | [9 empty] | 16
| 1 | 3 | 38 | < 100 | 78.2879 | 90 | 40 | 130 | [9 empty] | 16
| 1 | 4 | 35 | < 100 | 83.1789 | 80 | 40 | 140 | [9 empty] | 79.3256
| 1 | 5 | 37 | < 100 | 80.0767 | 110 | 40 | 150 | [9 empty] | 22
| 1 | 6 | 36 | < 100 | 79.474 | 90 | 30 | 160
| Mean |  | 37 |  | 80.2198 | 95 | 40 | 145
| Median |  | 37 |  | 79.9782 | 95 | 40 | 145
| Std Dev. |  | 1.4142 |  | 1.6262 | 10.4881 | 6.3246 | 10.4881
| Rel.Std.Dev. |  | 3.822% |  | 2.027% | 11.04% | 15.81% | 7.233%
| PDM3 |  | -53.36% |  | 1.127% | 19.76% | -49.57% | 82.79%
| Z-Score (Absolute) |  | 0.73 | NA | 0 | 0.26 | 0.67 | 1.11
| Analytical results for B in OREAS 461 (Indicative Value 151 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*OES | PF*OES | [13 empty] | 0
0
| 1 | 1 | 158 | 140 | [13 empty] | 1
| 1 | 2 | 157 | 140 | [13 empty] | 17
| 1 | 3 | 165 | 140 | [13 empty] | 16
| 1 | 4 | 161 | 160 | [13 empty] | 150.8333
| 1 | 5 | 176 | 140 | [13 empty] | 23
| 1 | 6 | 153 | 120
| Mean |  | 161.6667 | 140
| Median |  | 159.5 | 140
| Std Dev. |  | 8.0911 | 12.6491
| Rel.Std.Dev. |  | 5.005% | 9.035%
| PDM3 |  | 7.182% | -7.182%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for BaO in OREAS 461 (Certified Value 1037 ppm) | [16 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1090.8 | 1006 | 1093.1 | 1085.2 | 1035 | 953.1648 | 1075.2 | 1030.19 | 915.8248 | 1020.5 | 1061 | 1100.8424 | 1038 | 1105 | 893 | 1
| 1 | 2 | 1127.7 | 1060.7 | 1084.1 | 1054 | 1017.1 | 956.972 | 1064 | 1028.3 | 921.1304 | 1042.8 | 1061 | 1108.1789 | 1027 | 1083 | 916 | 2
| 1 | 3 | 1138.8 | 1010.4 | 1116.5 | 1085.2 | 1046.2 | 971.0543 | 1054 | 1010.1 | 922.3987 | 1010.4 | 1061 | 1066.1613 | 1038 | 1083 | 1027 | 16
| 1 | 4 | 1114.3 | 985.9 | 1081.9 | 1098.6 | 1062.9 | 959.6896 | 1029.4 | 1023.27 | 920.4605 | 1040.6 | 1061 | 1090.9579 | 1027 | 1083 | 938 | 1037.0501
| 1 | 5 | 1107.6 | 1055.1 | 1122.1 | 1103.1 | 1043.9 | 973.503 | 1093.1 | 1004.74 | 916.0034 | 1038.3 | 1061 | 1063.1553 | 994 | 1128 | 904 | 17
| 1 | 6 | 1089.7 | 1017.1 | 1088.6 | 1105.3 | 1056.2 | 995.6261 | 1030.5 | 1023.5 | 915.0655 | 1018.2 | 1061 | 1082.6852 | 1038 | 1139 | 927
| Mean |  | 1111.4833 | 1022.5333 | 1097.7167 | 1088.5667 | 1043.55 | 968.335 | 1057.7 | 1020.0167 | 918.4805 | 1028.4667 | 1061 | 1085.3302 | 1027 | 1103.5 | 934.1667
| Median |  | 1110.95 | 1013.75 | 1090.85 | 1091.9 | 1045.05 | 965.372 | 1059 | 1023.385 | 918.2319 | 1029.4 | 1061 | 1086.8215 | 1032.5 | 1094 | 921.5
| Std Dev. |  | 19.6698 | 29.3577 | 17.2474 | 19.0232 | 16.2129 | 15.575 | 25.1138 | 10.262 | 3.1983 | 13.7451 | 0 | 18.2238 | 17.0411 | 24.994 | 48.2055
| Rel.Std.Dev. |  | 1.77% | 2.871% | 1.571% | 1.748% | 1.554% | 1.608% | 2.374% | 1.006% | 0.3482% | 1.336% | 0% | 1.679% | 1.659% | 2.265% | 5.16%
| PDM3 |  | 7.177% | -1.4% | 5.85% | 4.968% | 0.6268% | -6.626% | 1.991% | -1.642% | -11.43% | -0.8277% | 2.309% | 4.656% | -0.9691% | 6.408% | -9.921%
| Z-Score (Absolute) |  | 1.1 | 0.34 | 0.87 | 0.73 | 0 | 1.21 | 0.23 | 0.38 | 2.02 | 0.24 | 0.28 | 0.67 | 0.16 | 0.97 | 2.07
| Analytical results for Be in OREAS 461 (Certified Value 2.3 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 17 | 22 | 23 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 2.4177 | 3 | 2.2 | 1.1375 | 2 | 2 | < 5 | [8 empty] | 1
| 1 | 2 | 2.1938 | 3 | 2.2 | 0.97849 | 2 | 2 | < 5 | [8 empty] | #N/A
| 1 | 3 | 2.2676 | 3 | 2.3 | 1.0965 | 2 | 2 | < 5 | [8 empty] | 16
| 1 | 4 | 2.2226 | 2 | 2.3 | 1.18 | 2 | 2 | < 5 | [8 empty] | 2.3019
| 1 | 5 | 2.1672 | 3 | 2.3 | 1.0649 | 2 | 3 | < 5 | [8 empty] | 18
| 1 | 6 | 2.2874 | 3 | 2.2 | 1.1239 | 2 | 2 | < 5
| Mean |  | 2.2594 | 2.8333 | 2.25 | 1.0969 | 2 | 2.1667
| Median |  | 2.2451 | 3 | 2.25 | 1.1102 | 2 | 2
| Std Dev. |  | 0.08949 | 0.40825 | 0.054772 | 0.069761 | 0 | 0.40825
| Rel.Std.Dev. |  | 3.961% | 14.41% | 2.434% | 6.36% | 0% | 18.84%
| PDM3 |  | -1.846% | 23.09% | -2.254% | -52.35% | -13.11% | -5.874%
| Z-Score (Absolute) |  | 0.27 | 3.25 | 0.22 | 5.78 | 1.08 | 0.22 | NA
| Analytical results for Bi in OREAS 461 (Certified Value 2.49 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 2.6178 | 2.6 | 2 | 1.2 | 3 | 2.4 | [9 empty] | 1
| 1 | 2 | 2.6915 | 2.6 | 2 | 1.1 | 2 | 2.2 | [9 empty] | #N/A
| 1 | 3 | 3.154 | 2.6 | 2 | 1.1 | 3 | 2.4 | [9 empty] | 16
| 1 | 4 | 3.0122 | 2.5 | 2 | 1.3 | 2 | 2.3 | [9 empty] | 2.4852
| 1 | 5 | 2.7663 | 2.6 | 2 | 1.1 | 3 | 2.1 | [9 empty] | 19
| 1 | 6 | 2.9142 | 2.6 | 2.5 | 1.2 | 3 | 2
| Mean |  | 2.8594 | 2.5833 | 2.0833 | 1.1667 | 2.6667 | 2.2333
| Median |  | 2.8403 | 2.6 | 2 | 1.15 | 3 | 2.25
| Std Dev. |  | 0.20416 | 0.040825 | 0.20412 | 0.08165 | 0.5164 | 0.1633
| Rel.Std.Dev. |  | 7.14% | 1.58% | 9.798% | 6.999% | 19.36% | 7.312%
| PDM3 |  | 15.06% | 3.949% | -16.17% | -53.06% | 7.302% | -10.13%
| Z-Score (Absolute) |  | 1.04 | 0.4 | 0.75 | 2.87 | 0.6 | 0.4
| Analytical results for CaO in OREAS 461 (Certified Value 1.76 wt.%) | [16 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.777 | 1.82 | 1.805 | 1.7689 | 1.7 | 1.875 | 1.7452 | 1.81 | 1.7994 | 1.96 | 1.68 |  |  |  |  | 1
| 1 | 2 | 1.7672 | 1.82 | 1.805 | 1.611 | 1.89 | 1.721 | 1.7503 | 1.84 | 1.7252 | 1.96 | 1.68 |  |  |  |  | 3
| 1 | 3 | 1.7378 | 1.82 | 1.791 | 1.6038 | 1.72 | 1.777 | 1.7563 | 1.8 | 1.6636 | 1.82 | 1.68 |  |  |  |  | 16
| 1 | 4 | 1.7588 | 1.82 | 1.791 | 1.614 | 1.73 | 1.735 | 1.7643 | 1.82 | 1.6874 | 1.82 | 1.68 |  |  |  |  | 1.7609
| 1 | 5 | 1.7644 | 1.96 | 1.791 | 1.6298 | 1.7 | 1.749 | 1.7282 | 1.77 | 1.728 | 1.96 | 1.68 |  |  |  |  | 20
| 1 | 6 | 1.7322 | 1.82 | 1.791 | 1.665 | 1.69 | 1.819 | 1.7553 | 1.82 | 1.8092 | 1.96 | 1.68
| Mean |  | 1.7562 | 1.8433 | 1.7957 | 1.6488 | 1.7383 | 1.7793 | 1.7499 | 1.81 | 1.7355 | 1.9133 | 1.68
| Median |  | 1.7616 | 1.82 | 1.791 | 1.6219 | 1.71 | 1.763 | 1.7528 | 1.815 | 1.7266 | 1.96 | 1.68
| Std Dev. |  | 0.017562 | 0.057155 | 0.0072296 | 0.062816 | 0.075741 | 0.058343 | 0.012374 | 0.023664 | 0.058584 | 0.072296 | 0
| Rel.Std.Dev. |  | 1% | 3.101% | 0.4026% | 3.81% | 4.357% | 3.279% | 0.7071% | 1.307% | 3.376% | 3.779% | 0%
| PDM3 |  | -0.2631% | 4.683% | 1.976% | -6.367% | -1.28% | 1.049% | -0.6218% | 2.79% | -1.442% | 8.659% | -4.592%
| Z-Score (Absolute) |  | 0 | 0.88 | 0.55 | 1.84 | 0.67 | 0.32 | 0.09 | 0.75 | 0.29 | 2.18 | 1.08
| Analytical results for Cd in OREAS 461 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | < 1 | < 1 | < 10 | < 0.2 | [11 empty] | 1
| 1 | 2 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | 18
| 1 | 3 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | 16
| 1 | 4 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | < 1
| 1 | 5 | < 1 | < 1 | < 10 | 0.2 | [11 empty] | 24
| 1 | 6 | < 1 | < 1 | < 10 | 0.4
| Mean |  |  |  |  | 0.3
| Median |  |  |  |  | 0.3
| Std Dev. |  |  |  |  | 0.070711
| Rel.Std.Dev. |  |  |  |  | 23.57%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for CeO2 in OREAS 461 (Certified Value 4311 ppm) | [16 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 4471.4 | 4152 | 4385.4 | 4324 | 4139.7 | 4045.6567 | 4504.5 | 3977.44 | 4051.1528 | 3482 | 4351 | 4960.3275 | 4471.38 | 4398 | 4238 | 1
| 1 | 2 | 4495.9 | 4520.5 | 4348.5 | 4201.1 | 4201.1 | 4033.6357 | 4477.5 | 4187.62 | 4002.2118 | 3453 | 4394 | 4987.4646 | 4299.4 | 4422 | 4238 | #N/A
| 1 | 3 | 4594.2 | 4238 | 4397.7 | 4348.5 | 4287.1 | 4100.955 | 4450.5 | 4250.26 | 4136.891 | 3396 | 4444.4 | 4747.5198 | 4373.1 | 4385 | 4631 | 16
| 1 | 4 | 4459.1 | 4139.7 | 4324 | 4385.4 | 4287.1 | 4083.1963 | 4435.8 | 4088.36 | 4052.7224 | 3347 | 4413.6 | 4845.5933 | 4483.66 | 4435 | 4250 | 4311.4355
| 1 | 5 | 4434.5 | 4471.4 | 4422.2 | 4348.5 | 4250.3 | 4087.2663 | 4430.8 | 4139.71 | 4129.3001 | 3417 | 4414.9 | 4810.3425 | 4557.36 | 4484 | 4263 | 21
| 1 | 6 | 4446.8 | 4336.3 | 4336.3 | 4373.1 | 4262.5 | 4133.6237 | 4398.9 | 3966.14 | 4099.3341 | 3330 | 4573.3 | 4875.4343 | 4520.51 | 4606 | 4361
| Mean |  | 4483.65 | 4309.65 | 4369.0167 | 4330.1 | 4237.9667 | 4080.7223 | 4449.6667 | 4101.5883 | 4078.602 | 3404.1667 | 4431.8667 | 4871.1137 | 4450.9017 | 4455 | 4330.1667
| Median |  | 4465.25 | 4287.15 | 4366.95 | 4348.5 | 4256.4 | 4085.2313 | 4443.15 | 4114.035 | 4076.0282 | 3406.5 | 4414.25 | 4860.5138 | 4477.52 | 4428.5 | 4256.5
| Std Dev. |  | 58.1339 | 161.3699 | 38.5679 | 66.7269 | 57.6154 | 36.6159 | 37.1687 | 113.9228 | 52.2616 | 59.0641 | 75.8358 | 90.703 | 96.6199 | 81.5843 | 154.5179
| Rel.Std.Dev. |  | 1.297% | 3.744% | 0.8828% | 1.541% | 1.36% | 0.8973% | 0.8353% | 2.778% | 1.281% | 1.735% | 1.711% | 1.862% | 2.171% | 1.831% | 3.568%
| PDM3 |  | 3.994% | -0.04141% | 1.336% | 0.4329% | -1.704% | -5.351% | 3.206% | -4.867% | -5.4% | -21.04% | 2.793% | 12.98% | 3.235% | 3.33% | 0.4345%
| Z-Score (Absolute) |  | 0.86 | 0.11 | 0.22 | 0 | 0.51 | 1.39 | 0.67 | 1.28 | 1.4 | 5.17 | 0.57 | 3.02 | 0.67 | 0.7 | 0.34
| Analytical results for Co in OREAS 461 (Certified Value 11.5 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 13.5 | 13.9258 | 9.9 | < 50 | 9.5234 | 9 | < 10 | 20 | 11.6 | [6 empty] | 1
| 1 | 2 | 13.5 | 13.8322 | 9.7 | < 50 | 9.4401 | 11.1 | < 10 | 20 | 10.9 | [6 empty] | #N/A
| 1 | 3 | 13.7 | 13.831 | 9.7 | < 50 | 9.4924 | 10.8 | < 10 | 20 | 13.8 | [6 empty] | 16
| 1 | 4 | 13.3 | 14.0516 | 9.9 | < 50 | 10.8107 | 6.9 | < 10 | 20 | 10.6 | [6 empty] | 11.5257
| 1 | 5 | 13.3 | 13.7931 | 9.9 | < 50 | 9.2972 | 11.5 | < 10 | 20 | 11.5 | [6 empty] | 22
| 1 | 6 | 13.1 | 13.8425 | 9.9 | < 50 | 9.7244 | 10.9 | < 10 | 20 | 11.6
| Mean |  | 13.4 | 13.8794 | 9.8333 |  | 9.7147 | 10.0333 |  | 20 | 11.6667
| Median |  | 13.4 | 13.8373 | 9.9 |  | 9.5079 | 10.85 |  | 20 | 11.55
| Std Dev. |  | 0.20976 | 0.095082 | 0.10328 |  | 0.55443 | 1.7614 |  | 0 | 1.1237
| Rel.Std.Dev. |  | 1.565% | 0.6851% | 1.05% |  | 5.707% | 17.56% |  | 0% | 9.632%
| PDM3 |  | 16.26% | 20.42% | -14.68% |  | -15.71% | -12.95% |  | 73.53% | 1.223%
| Z-Score (Absolute) |  | 0.64 | 0.81 | 0.67 | NA | 0.72 | 0.37 | NA | 3.07 | 0
| Analytical results for Cr2O3 in OREAS 461 (Certified Value 875 ppm) | [16 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 0
0
| 1 | 1 | 950 | 833 | 906 | 979 | 804 | < 1500 | 847.7 | 877.7652 | 900 | 877 | 827.0195 | 880 | 1096 | 658 | 892 | 1
| 1 | 2 | 965 | 892 | 892 | 965 | 789 | < 1500 | 834.5 | 875.3318 | 900 | 877 | 856.0046 | 880 | 1023 | 658 | 892 | 4
| 1 | 3 | 979 | 848 | 877 | 979 | 818 | < 1500 | 821.4 | 866.348 | 900 | 877 | 840.5187 | 730 | 1125 | 687 | 965 | 16
| 1 | 4 | 965 | 833 | 906 | 979 | 804 | < 1500 | 838.9 | 868.8881 | 900 | 877 | 840.6688 | 730 | 1125 | 658 | 877 | 875.4856
| 1 | 5 | 950 | 892 | 906 | 979 | 789 | < 1500 | 828.7 | 867.4426 | 900 | 877 | 828.3102 | 730 | 848 | 658 | 892 | 23
| 1 | 6 | 950 | 848 | 906 | 994 | 804 | < 1500 | 817 | 871.5597 | 900 | 877 | 823.0455 | 880 | 1067 | 687 | 877
| Mean |  | 959.8333 | 857.6667 | 898.8333 | 979.1667 | 801.3333 |  | 831.3667 | 871.2226 | 900 | 877 | 835.9279 | 805 | 1047.3333 | 667.6667 | 899.1667
| Median |  | 957.5 | 848 | 906 | 979 | 804 |  | 831.6 | 870.2239 | 900 | 877 | 834.4144 | 805 | 1081.5 | 658 | 892
| Std Dev. |  | 11.9234 | 27.4275 | 12.0734 | 9.1742 | 10.9848 |  | 11.3683 | 4.5454 | 0 | 0 | 12.2472 | 82.1584 | 105.0194 | 14.9755 | 33.0782
| Rel.Std.Dev. |  | 1.242% | 3.198% | 1.343% | 0.9369% | 1.371% |  | 1.367% | 0.5217% | 0% | 0% | 1.465% | 10.21% | 10.03% | 2.243% | 3.679%
| PDM3 |  | 9.634% | -2.035% | 2.667% | 11.84% | -8.47% |  | -5.039% | -0.4869% | 2.8% | 0.173% | -4.518% | -8.051% | 19.63% | -23.74% | 2.705%
| Z-Score (Absolute) |  | 1.43 | 0.28 | 0.49 | 1.75 | 1.21 | NA | 0.71 | 0.05 | 0.43 | 0.05 | 0.64 | 1.17 | 3.56 | 3.45 | 0.19
| Analytical results for Cs in OREAS 461 (Certified Value 0.79 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.87 | 0.76 | 0.83 | 0.77 | 0.74 | 0.77175 | 0.9 | 0.8 | 0.57 | < 1 | 0.57329 | 0.3 | 1 | 0.8 |  | 1
| 1 | 2 | 0.84 | 0.8 | 0.8 | 0.76 | 0.72 | 0.80135 | 0.8 | 0.8 | 0.55 | < 1 | 0.54669 | 0.3 | < 1 | 0.9 |  | #N/A
| 1 | 3 | 0.9 | 0.59 | 0.76 | 0.77 | 0.74 | 0.8125 | 0.8 | 0.8 | 0.52 | < 1 | 0.51027 | 0.3 | < 1 | 0.9 |  | 16
| 1 | 4 | 0.79 | 0.76 | 0.79 | 0.75 | 0.75 | 0.78179 | 0.7 | 0.8 | 0.53 | < 1 | 0.54631 | 0.2 | < 1 | 1.1 |  | 0.78598
| 1 | 5 | 0.79 | 0.71 | 0.82 | 0.69 | 0.79 | 0.77057 | 0.7 | 0.8 | 0.52 | < 1 | 0.52435 | 0.3 | < 1 | 1.1 |  | 24
| 1 | 6 | 0.8 | 0.64 | 0.79 | 0.76 | 0.73 | 0.77685 | 0.8 | 0.8 | 0.5 | < 1 | 0.5113 | 0.3 | < 1 | 1.1
| Mean |  | 0.83167 | 0.71 | 0.79833 | 0.75 | 0.745 | 0.7858 | 0.78333 | 0.8 | 0.53167 |  | 0.53537 | 0.28333 | 1 | 0.98333
| Median |  | 0.82 | 0.735 | 0.795 | 0.76 | 0.74 | 0.77932 | 0.8 | 0.8 | 0.525 |  | 0.53533 | 0.3 | 1 | 1
| Std Dev. |  | 0.046224 | 0.080498 | 0.024833 | 0.030332 | 0.02429 | 0.017205 | 0.075277 | 1.2162e-16 | 0.024833 |  | 0.024563 | 0.040825 |  | 0.13292
| Rel.Std.Dev. |  | 5.558% | 11.34% | 3.111% | 4.044% | 3.26% | 2.19% | 9.61% | 1.52e-14% | 4.671% |  | 4.588% | 14.41% |  | 13.52%
| PDM3 |  | 5.813% | -9.667% | 1.572% | -4.577% | -5.214% | -0.02273% | -0.3364% | 1.784% | -32.36% |  | -31.89% | -63.95% | 27.23% | 25.11%
| Z-Score (Absolute) |  | 0.67 | 1.02 | 0.21 | 0.3 | 0.66 | 0.03 | 0 | 0.23 | 3.51 | NA | 3.46 | 6.98 | 3.02 | 1.16
| Analytical results for Cu in OREAS 461 (Certified Value 60 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | 58 | < 100 | 69.7187 | 66 | < 80 | 50 | 60 | [8 empty] | 1
| 1 | 2 | 58 | < 100 | 71.7051 | 476 | < 80 | 50 | 60 | [8 empty] | #N/A
| 1 | 3 | 57 | < 100 | 76.0173 | 19 | < 80 | 100 | 60 | [8 empty] | 16
| 1 | 4 | 70 | < 100 | 70.4566 | 29 | < 80 | 50 | 60 | [8 empty] | 60.1422
| 1 | 5 | 58 | < 100 | 68.527 | 24 | < 80 | 100 | 60 | [8 empty] | 25
| 1 | 6 | 73 | < 100 | 69.0418 | 35 | < 80 | 100 | 60
| Mean |  | 62.3333 |  | 70.9111 | 108.1667 |  | 75 | 60
| Median |  | 58 |  | 70.0877 | 32 |  | 75 | 60
| Std Dev. |  | 7.174 |  | 2.739 | 180.9612 |  | 27.3861 | 0
| Rel.Std.Dev. |  | 11.51% |  | 3.863% | 167.3% |  | 36.51% | 0%
| PDM3 |  | 3.643% |  | 17.91% | 79.85% |  | 24.7% | -0.2365%
| Z-Score (Absolute) |  | 0 | NA | 0.67 | 1.61 | NA | 0.93 | 0.01
| Analytical results for Dy2O3 in OREAS 461 (Certified Value 40 ppm) | [16 empty] | Certified
Dy2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 37.185 | 39.481 | 39.137 | 40.284 | 37.759 | 40.6336 | 42.69 | 40.05 | 34.3777 | 32.17 | 40.17 | 38.5463 | 65.648 | 48.78 | 38.1 | 1
| 1 | 2 | 37.185 | 42.924 | 40.514 | 39.366 | 37.759 | 39.8913 | 42.24 | 40.05 | 34.7431 | 32.457 | 40.17 | 39.0687 | 62.664 | 40.17 | 38.68 | 5
| 1 | 3 | 38.792 | 41.317 | 39.366 | 40.17 | 38.448 | 41.8137 | 41.43 | 39.71 | 31.0506 | 32.009 | 41.32 | 37.2362 | 62.55 | 41.32 | 42.24 | 16
| 1 | 4 | 37.53 | 38.218 | 41.088 | 41.202 | 39.366 | 41.3917 | 40.74 | 39.94 | 32.4196 | 31.527 | 41.32 | 38.2401 | 54.057 | 43.61 | 40.28 | 39.9771
| 1 | 5 | 37.874 | 42.35 | 40.284 | 40.743 | 39.022 | 40.7441 | 42.35 | 39.02 | 34.5048 | 31.803 | 40.17 | 36.6809 | 67.37 | 42.46 | 38.79 | 26
| 1 | 6 | 37.3 | 39.137 | 38.448 | 40.514 | 39.251 | 40.5381 | 41.09 | 39.71 | 33.7906 | 31.367 | 42.46 | 37.184 | 64.156 | 43.04 | 40.28
| Mean |  | 37.6443 | 40.5712 | 39.8062 | 40.3798 | 38.6008 | 40.8354 | 41.7567 | 39.7467 | 33.481 | 31.8888 | 40.935 | 37.826 | 62.7408 | 43.23 | 39.7283
| Median |  | 37.415 | 40.399 | 39.825 | 40.399 | 38.735 | 40.6888 | 41.835 | 39.825 | 34.0841 | 31.906 | 40.745 | 37.7382 | 63.41 | 42.75 | 39.535
| Std Dev. |  | 0.62024 | 1.8994 | 0.98576 | 0.61756 | 0.72475 | 0.67754 | 0.77994 | 0.38785 | 1.4543 | 0.40667 | 0.93571 | 0.92799 | 4.6332 | 2.9862 | 1.5197
| Rel.Std.Dev. |  | 1.648% | 4.682% | 2.476% | 1.529% | 1.878% | 1.659% | 1.868% | 0.9758% | 4.344% | 1.275% | 2.286% | 2.453% | 7.385% | 6.908% | 3.825%
| PDM3 |  | -5.835% | 1.486% | -0.4276% | 1.007% | -3.443% | 2.147% | 4.451% | -0.5764% | -16.25% | -20.23% | 2.396% | -5.381% | 56.94% | 8.137% | -0.6223%
| Z-Score (Absolute) |  | 1.34 | 0.43 | 0 | 0.32 | 0.67 | 0.58 | 1.09 | 0.03 | 3.54 | 4.43 | 0.63 | 1.11 | 13.81 | 1.3 | 0.04
| Analytical results for Er2O3 in OREAS 461 (Certified Value 10.1 ppm) | [16 empty] | Certified
Er2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 9.525 | 10.166 | 10.143 | 9.948 | 9.663 | 10.0823 | 10.41 | 9.95 | 9.8787 | 8.199 | 10.06 | 9.2284 | 8.05 | 11.43 | 10.349 | 1
| 1 | 2 | 9.411 | 10.555 | 10.337 | 10.337 | 10.006 | 10.5492 | 10.63 | 9.95 | 9.572 | 8.004 | 10.18 | 9.2546 | 7.65 | 10.86 | 10.452 | 6
| 1 | 3 | 9.983 | 10.497 | 10.497 | 9.903 | 10.166 | 10.166 | 10.52 | 9.83 | 9.6057 | 7.787 | 10.29 | 9.025 | 7.604 | 10.86 | 10.166 | 16
| 1 | 4 | 9.811 | 10.097 | 10.669 | 10.017 | 10.303 | 9.6639 | 10.18 | 9.72 | 9.8946 | 7.97 | 9.95 | 9.1291 | 6.621 | 10.86 | 10.303 | 10.0658
| 1 | 5 | 9.525 | 10.394 | 10.349 | 10.28 | 10.257 | 9.4385 | 10.63 | 9.72 | 9.6633 | 7.902 | 9.72 | 8.8324 | 8.119 | 10.86 | 10.943 | 27
| 1 | 6 | 9.308 | 10.189 | 10.017 | 10.417 | 10.006 | 9.9667 | 10.06 | 9.72 | 9.4623 | 7.81 | 10.29 | 8.7857 | 7.799 | 11.43 | 10.234
| Mean |  | 9.5938 | 10.3163 | 10.3353 | 10.1503 | 10.0668 | 9.9778 | 10.405 | 9.815 | 9.6794 | 7.9453 | 10.0817 | 9.0425 | 7.6405 | 11.05 | 10.4078
| Median |  | 9.525 | 10.2915 | 10.343 | 10.1485 | 10.086 | 10.0245 | 10.465 | 9.775 | 9.6345 | 7.936 | 10.12 | 9.0771 | 7.7245 | 10.86 | 10.326
| Std Dev. |  | 0.25425 | 0.19108 | 0.23494 | 0.2203 | 0.23335 | 0.39046 | 0.23839 | 0.11292 | 0.17341 | 0.15077 | 0.2214 | 0.1987 | 0.54085 | 0.29435 | 0.27982
| Rel.Std.Dev. |  | 2.65% | 1.852% | 2.273% | 2.17% | 2.318% | 3.913% | 2.291% | 1.15% | 1.792% | 1.898% | 2.196% | 2.197% | 7.079% | 2.664% | 2.689%
| PDM3 |  | -4.688% | 2.489% | 2.678% | 0.8403% | 0.01075% | -0.8741% | 3.37% | -2.491% | -3.838% | -21.07% | 0.1581% | -10.17% | -24.09% | 9.778% | 3.398%
| Z-Score (Absolute) |  | 1.19 | 0.63 | 0.67 | 0.21 | 0 | 0.22 | 0.85 | 0.63 | 0.97 | 5.32 | 0.04 | 2.57 | 5.58 | 2.48 | 0.59
| Analytical results for Eu2O3 in OREAS 461 (Certified Value 54 ppm) | [16 empty] | Certified
Eu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 50.83 | 54.31 | 51.87 | 53.73 | 50.6 | 51.5842 | 59.28 | 51.06 | 33.2965 | 45.204 | 51.53 | 56.6333 | 101.895 | 61.37 | 53.96 | 1
| 1 | 2 | 51.06 | 57.43 | 52.92 | 51.64 | 52.45 | 53.2888 | 57.66 | 51.87 | 33.4512 | 44.845 | 52.68 | 57.3781 | 96.453 | 55 | 54.42 | 7
| 1 | 3 | 53.15 | 55.93 | 52.57 | 52.34 | 52.68 | 52.335 | 58.59 | 51.64 | 32.7417 | 44.394 | 53.26 | 54.5928 | 96.221 | 57.32 | 60.91 | 16
| 1 | 4 | 51.87 | 52.22 | 54.07 | 54.19 | 53.38 | 51.7287 | 58.47 | 52.57 | 34.1768 | 43.896 | 53.26 | 55.9721 | 83.137 | 56.16 | 52.11 | 54.0907
| 1 | 5 | 51.06 | 57.08 | 53.73 | 53.38 | 53.49 | 52.1825 | 59.28 | 51.06 | 34.2341 | 44.579 | 53.26 | 54.9792 | 102.127 | 58.47 | 54.88 | 28
| 1 | 6 | 51.06 | 53.03 | 53.15 | 54.31 | 53.26 | 52.9843 | 57.55 | 50.83 | 34.8675 | 44.081 | 55 | 56.6661 | 98.885 | 59.05 | 55.58
| Mean |  | 51.505 | 55 | 53.0517 | 53.265 | 52.6433 | 52.3506 | 58.4717 | 51.505 | 33.7946 | 44.4998 | 53.165 | 56.037 | 96.453 | 57.895 | 55.31
| Median |  | 51.06 | 55.12 | 53.035 | 53.555 | 52.97 | 52.2588 | 58.53 | 51.35 | 33.814 | 44.4865 | 53.26 | 56.3027 | 97.669 | 57.895 | 54.65
| Std Dev. |  | 0.88192 | 2.1532 | 0.79404 | 1.0652 | 1.082 | 0.67593 | 0.75221 | 0.65381 | 0.77006 | 0.48425 | 1.1227 | 1.0732 | 7.0013 | 2.2566 | 2.9824
| Rel.Std.Dev. |  | 1.712% | 3.915% | 1.497% | 2% | 2.055% | 1.291% | 1.286% | 1.269% | 2.279% | 1.088% | 2.112% | 1.915% | 7.259% | 3.898% | 5.392%
| PDM3 |  | -4.78% | 1.681% | -1.921% | -1.526% | -2.676% | -3.217% | 8.099% | -4.78% | -37.52% | -17.73% | -1.711% | 3.598% | 78.32% | 7.033% | 2.254%
| Z-Score (Absolute) |  | 0.67 | 0.74 | 0.05 | 0.04 | 0.21 | 0.33 | 2.16 | 0.67 | 7.87 | 3.52 | 0 | 1.17 | 18.67 | 1.92 | 0.42
| Analytical results for Fe2O3 in OREAS 461 (Certified Value 45.43 wt.%) | [16 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 43.3158 | 47.838 | 45.279 | 46.6234 | 43.14 | 47.237 | 45.5029 | 43.3 | 45.636 | 47.18 | > 35.7 |  |  |  |  | 1
| 1 | 2 | 43.495 | 47.852 | 45.593 | 48.0319 | 43.97 | 46.48 | 45.4059 | 42.2 | 44.935 | 47.32 | > 35.7 |  |  |  |  | #N/A
| 1 | 3 | 43.4123 | 46.236 | 45.321 | 47.618 | 42.8 | 47.094 | 45.5416 | 39.1 | 43.377 | 46.75 | > 35.7 |  |  |  |  | 16
| 1 | 4 | 43.1504 | 47.838 | 45.464 | 47.0693 | 43.73 | 47.266 | 45.9587 | 33.1 | 43.735 | 47.61 | > 35.7 |  |  |  |  | 45.4263
| 1 | 5 | 43.4399 | 46.98 | 45.865 | 47.1003 | 42.67 | 46.951 | 45.5319 | 41.9 | 45.422 | 48.47 | > 35.7 |  |  |  |  | 29
| 1 | 6 | 43.1642 | 46.265 | 45.307 | 46.8092 | 42.73 | 47.266 | 46.2108 | 42.6 | 46.408 | 49.18 | > 35.7
| Mean |  | 43.3296 | 47.1682 | 45.4715 | 47.2087 | 43.1733 | 47.049 | 45.692 | 40.3667 | 44.9188 | 47.7517
| Median |  | 43.3641 | 47.409 | 45.3925 | 47.0848 | 42.97 | 47.1655 | 45.5368 | 42.05 | 45.1785 | 47.465
| Std Dev. |  | 0.14564 | 0.7855 | 0.22646 | 0.52509 | 0.55406 | 0.30493 | 0.31816 | 3.8396 | 1.1631 | 0.90473
| Rel.Std.Dev. |  | 0.3361% | 1.665% | 0.498% | 1.112% | 1.283% | 0.6481% | 0.6963% | 9.512% | 2.589% | 1.895%
| PDM3 |  | -4.616% | 3.834% | 0.09946% | 3.924% | -4.96% | 3.572% | 0.5848% | -11.14% | -1.117% | 5.119%
| Z-Score (Absolute) |  | 0.95 | 0.67 | 0.05 | 0.68 | 1.01 | 0.62 | 0.05 | 1.29 | 0.28 | 0.91 | NA
| Analytical results for Ga in OREAS 461 (Certified Value 50 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 12 | 14 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 51.8 | 47.4 | 47.1 | 44.3 | 29.3579 | 14.3 | 58 | 44 | [7 empty] | 1
| 1 | 2 | 55.9 | 47.6 | 46.1 | 43.9 | 28.3775 | 14.3 | 60 | 43 | [7 empty] | #N/A
| 1 | 3 | 52.1 | 48 | 47.9 | 45.5 | 28.9499 | 15 | 62 | 60 | [7 empty] | 16
| 1 | 4 | 51.8 | 48.2 | 48.3 | 45.1 | 29.6587 | 16.8 | 62 | 45 | [7 empty] | 50.0389
| 1 | 5 | 56 | 50.2 | 47.5 | 45.7 | 29.7931 | 15.9 | 62 | 46 | [7 empty] | 30
| 1 | 6 | 51.1 | 48.3 | 48.2 | 45 | 29.3417 | 16.5 | 62 | 46
| Mean |  | 53.1167 | 48.2833 | 47.5167 | 44.9167 | 29.2465 | 15.4667 | 61 | 47.3333
| Median |  | 51.95 | 48.1 | 47.7 | 45.05 | 29.3498 | 15.45 | 62 | 45.5
| Std Dev. |  | 2.2194 | 1.0008 | 0.82563 | 0.69402 | 0.5164 | 1.093 | 1.6733 | 6.314
| Rel.Std.Dev. |  | 4.178% | 2.073% | 1.738% | 1.545% | 1.766% | 7.067% | 2.743% | 13.34%
| PDM3 |  | 6.151% | -3.508% | -5.041% | -10.24% | -41.55% | -69.09% | 21.91% | -5.407%
| Z-Score (Absolute) |  | 1.04 | 0.31 | 0.2 | 0.2 | 2.55 | 4.62 | 2.31 | 0.21
| Analytical results for Gd2O3 in OREAS 461 (Certified Value 116 ppm) | [16 empty] | Certified
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 115.84 | 113.99 | 117.57 | 105.23 | 105 | 119.64 | 107.71 | 124.9611 | 100.53 | 116.99 | 136 | 147.533 | 131.4 | 111.92 |  | 1
| 1 | 2 | 115.84 | 123.9 | 112.61 | 103.62 | 105.58 | 115.14 | 108.402 | 122.2499 | 100.68 | 114.11 | 136 | 140.617 | 119.9 | 110.88 |  | 8
| 1 | 3 | 121.02 | 123.33 | 111.69 | 104.89 | 109.96 | 115.03 | 107.03 | 124.6833 | 99.089 | 117.57 | 139 | 147.533 | 122.2 | 119.87 |  | 16
| 1 | 4 | 115.84 | 111 | 115.26 | 107.88 | 109.84 | 118.72 | 106.512 | 125.6933 | 99.193 | 118.14 | 142 | 152.143 | 122.2 | 111.69 |  | 115.7064
| 1 | 5 | 118.14 | 119.87 | 115.84 | 108.11 | 110.19 | 117.22 | 107.826 | 125.6008 | 102.916 | 118.14 | 137 | 153.296 | 122.2 | 114.45 |  | 31
| 1 | 6 | 115.84 | 113.99 | 110.3 | 108.11 | 110.19 | 114.8 | 106.823 | 123.4971 | 97.764 | 120.45 | 137 | 149.838 | 129.1 | 114.57
| Mean |  | 117.0867 | 117.68 | 113.8783 | 106.3067 | 108.46 | 116.7583 | 107.3838 | 124.4476 | 100.0287 | 117.5667 | 137.8333 | 148.4933 | 124.5 | 113.8967
| Median |  | 115.84 | 116.93 | 113.935 | 106.555 | 109.9 | 116.18 | 107.37 | 124.8222 | 99.8615 | 117.855 | 137 | 148.6855 | 122.2 | 113.185
| Std Dev. |  | 2.1353 | 5.427 | 2.778 | 1.9679 | 2.466 | 2.0884 | 0.71251 | 1.3367 | 1.7733 | 2.0617 | 2.3166 | 4.5183 | 4.6 | 3.296
| Rel.Std.Dev. |  | 1.824% | 4.612% | 2.439% | 1.851% | 2.274% | 1.789% | 0.6635% | 1.074% | 1.773% | 1.754% | 1.681% | 3.043% | 3.695% | 2.894%
| PDM3 |  | 1.193% | 1.706% | -1.58% | -8.124% | -6.263% | 0.9092% | -7.193% | 7.555% | -13.55% | 1.608% | 19.12% | 28.34% | 7.6% | -1.564%
| Z-Score (Absolute) |  | 0.02 | 0.1 | 0.23 | 0.87 | 0.69 | 0.02 | 0.78 | 0.67 | 1.4 | 0.09 | 1.74 | 2.72 | 0.68 | 0.22
| Analytical results for Ge in OREAS 461 (Indicative Value 2.5 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 20 | 22 | 23 | [12 empty] | ppm
|  |  | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 1 | < 20 | 3 | [12 empty] | 1
| 1 | 2 | < 1 | < 20 | 2 | [12 empty] | 20
| 1 | 3 | < 1 | < 20 | 3 | [12 empty] | 16
| 1 | 4 | < 1 | < 20 | 2 | [12 empty] | 2.5
| 1 | 5 | < 1 | < 20 | 2 | [12 empty] | 21
| 1 | 6 | < 1 | < 20 | 3
| Mean |  |  |  | 2.5
| Median |  |  |  | 2.5
| Std Dev. |  |  |  | 0.54772
| Rel.Std.Dev. |  |  |  | 21.91%
| PDM3 |  |  |  | 0%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for HfO2 in OREAS 461 (Certified Value 16.6 ppm) | [16 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 20 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 15.92 | 17.57 | 15.45 | 16.75 | 16.51 | 16.98 | 13.68 | 13.456 | 15.3 | 17.7 | 16.5 |  |  |  |  | 1
| 1 | 2 | 16.39 | 20.05 | 15.57 | 16.27 | 16.63 | 16.98 | 13.44 | 13.833 | 16.5 | 16.5 | 15.3 |  |  |  |  | 9
| 1 | 3 | 16.63 | 17.81 | 15.21 | 16.27 | 16.27 | 16.98 | 13.44 | 12.984 | 16.5 | 16.5 | 17.7 |  |  |  |  | 16
| 1 | 4 | 16.39 | 17.22 | 16.27 | 17.22 | 17.1 | 16.86 | 13.92 | 13.137 | 16.5 | 17.7 | 16.5 |  |  |  |  | 16.6198
| 1 | 5 | 15.8 | 18.28 | 16.39 | 16.27 | 16.63 | 16.75 | 13.44 | 13.067 | 16.5 | 17.7 | 16.5 |  |  |  |  | 32
| 1 | 6 | 16.16 | 17.69 | 15.45 | 16.86 | 16.27 | 16.39 | 13.8 | 12.677 | 16.5 | 17.7 | 15.3
| Mean |  | 16.215 | 18.1033 | 15.7233 | 16.6067 | 16.5683 | 16.8233 | 13.62 | 13.1923 | 16.3 | 17.3 | 16.3
| Median |  | 16.275 | 17.75 | 15.51 | 16.51 | 16.57 | 16.92 | 13.56 | 13.102 | 16.5 | 17.7 | 16.5
| Std Dev. |  | 0.31488 | 1.0141 | 0.48574 | 0.40023 | 0.30714 | 0.23157 | 0.21128 | 0.40192 | 0.4899 | 0.61968 | 0.90333
| Rel.Std.Dev. |  | 1.942% | 5.602% | 3.089% | 2.41% | 1.854% | 1.377% | 1.551% | 3.047% | 3.006% | 3.582% | 5.542%
| PDM3 |  | -2.436% | 8.926% | -5.394% | -0.07921% | -0.3099% | 1.224% | -18.05% | -20.62% | -1.924% | 4.093% | -1.924%
| Z-Score (Absolute) |  | 0.13 | 1.85 | 0.78 | 0.39 | 0.34 | 0.67 | 3.55 | 4.12 | 0 | 1.3 | 0
| Analytical results for Ho2O3 in OREAS 461 (Certified Value 5.22 ppm) | [16 empty] | Certified
Ho2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.811 | 5.143 | 5.178 | 5.201 | 4.914 | 5.1091 | 5.5 | 5.15 | 3.2063 | 4.227 | 5.04 | 5.0075 | 5.705 | 6.41 | 4.845 | 1
| 1 | 2 | 4.685 | 5.498 | 5.407 | 5.029 | 5.017 | 5.4463 | 5.5 | 5.15 | 3.7663 | 4.147 | 5.27 | 4.873 | 5.475 | 5.5 | 4.742 | 10
| 1 | 3 | 5.006 | 5.372 | 5.292 | 5.258 | 5.04 | 5.423 | 5.61 | 5.04 | 3.4007 | 4.067 | 5.15 | 4.7664 | 5.601 | 5.73 | 5.567 | 16
| 1 | 4 | 4.857 | 5.109 | 5.453 | 5.327 | 4.994 | 5.557 | 5.38 | 5.15 | 3.4503 | 4.044 | 4.93 | 4.8545 | 5.911 | 5.96 | 4.8 | 5.2229
| 1 | 5 | 4.868 | 5.567 | 5.327 | 5.304 | 5.063 | 5.4854 | 5.5 | 5.04 | 3.6236 | 4.193 | 5.15 | 4.5666 | 5.888 | 5.96 | 4.479 | 33
| 1 | 6 | 4.719 | 5.109 | 5.281 | 5.372 | 5.097 | 5.2355 | 5.27 | 5.04 | 3.5919 | 4.101 | 5.38 | 4.7713 | 5.728 | 6.19 | 4.651
| Mean |  | 4.8243 | 5.2997 | 5.323 | 5.2485 | 5.0208 | 5.3761 | 5.46 | 5.095 | 3.5065 | 4.1298 | 5.1533 | 4.8065 | 5.718 | 5.9583 | 4.8473
| Median |  | 4.834 | 5.2575 | 5.3095 | 5.281 | 5.0285 | 5.4346 | 5.5 | 5.095 | 3.5211 | 4.124 | 5.15 | 4.8129 | 5.7165 | 5.96 | 4.771
| Std Dev. |  | 0.11547 | 0.20654 | 0.097695 | 0.12249 | 0.063395 | 0.16901 | 0.11815 | 0.060249 | 0.19646 | 0.071934 | 0.15983 | 0.14667 | 0.16683 | 0.32245 | 0.37563
| Rel.Std.Dev. |  | 2.393% | 3.897% | 1.835% | 2.334% | 1.263% | 3.144% | 2.164% | 1.183% | 5.603% | 1.742% | 3.102% | 3.052% | 2.918% | 5.412% | 7.749%
| PDM3 |  | -7.631% | 1.47% | 1.916% | 0.4899% | -3.869% | 2.932% | 4.539% | -2.449% | -32.86% | -20.93% | -1.332% | -7.972% | 9.479% | 14.08% | -7.191%
| Z-Score (Absolute) |  | 0.73 | 0.32 | 0.37 | 0.21 | 0.3 | 0.49 | 0.67 | 0.13 | 3.64 | 2.26 | 0 | 0.77 | 1.24 | 1.57 | 1
| Analytical results for In in OREAS 461 (Certified Value 0.6 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.63352 | 0.6 | 0.5 | 0.4 | 0.6 | 0.6 | [9 empty] | 1
| 1 | 2 | 0.501 | 0.6 | 0.5 | 0.3 | 0.6 | 0.7 | [9 empty] | #N/A
| 1 | 3 | 0.61738 | 0.6 | 0.5 | 0.3 | 0.6 | 0.6 | [9 empty] | 16
| 1 | 4 | 0.64719 | 0.6 | 0.5 | 0.5 | 0.6 | 0.6 | [9 empty] | 0.60113
| 1 | 5 | 0.69894 | 0.7 | 0.5 | 0.4 | 0.6 | 0.7 | [9 empty] | 34
| 1 | 6 | 0.53581 | 0.7 | 0.5 | 0.3 | 0.6 | 0.8
| Mean |  | 0.60564 | 0.63333 | 0.5 | 0.36667 | 0.6 | 0.66667
| Median |  | 0.62545 | 0.6 | 0.5 | 0.35 | 0.6 | 0.65
| Std Dev. |  | 0.073721 | 0.05164 | 0 | 0.08165 | 0 | 0.08165
| Rel.Std.Dev. |  | 12.17% | 8.154% | 0% | 22.27% | 0% | 12.25%
| PDM3 |  | 0.7505% | 5.358% | -16.82% | -39% | -0.1876% | 10.9%
| Z-Score (Absolute) |  | 0.04 | 0.44 | 1.47 | 3.38 | 0.04 | 0.91
| Analytical results for K2O in OREAS 461 (Certified Value 0.285 wt.%) | [16 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.359 | 0.325 | 0.277 | 0.34168 | 0.28 | 0.24 | 0.2388 | 0.3 | 0.289 | 0.24 | [5 empty] | 1
| 1 | 2 | 0.3385 | 0.301 | 0.277 | 0.33028 | 0.28 | 0.24 | 0.2398 | 0.3 | 0.277 | 0.24 | [5 empty] | 11
| 1 | 3 | 0.3385 | 0.313 | 0.265 | 0.34827 | 0.28 | 0.24 | 0.239 | 0.3 | 0.241 | 0.24 | [5 empty] | 16
| 1 | 4 | 0.3554 | 0.337 | 0.277 | 0.34548 | 0.29 | 0.24 | 0.2423 | 0.3 | 0.241 | 0.24 | [5 empty] | 0.28495
| 1 | 5 | 0.3385 | 0.325 | 0.277 | 0.33753 | 0.27 | 0.24 | 0.2349 | 0.28 | 0.277 | 0.24 | [5 empty] | 35
| 1 | 6 | 0.3433 | 0.325 | 0.277 | 0.33791 | 0.27 | 0.24 | 0.235 | 0.3 | 0.289 | 0.24
| Mean |  | 0.34553 | 0.321 | 0.275 | 0.34019 | 0.27833 | 0.24 | 0.2383 | 0.29667 | 0.269 | 0.24
| Median |  | 0.3409 | 0.325 | 0.277 | 0.3398 | 0.28 | 0.24 | 0.2389 | 0.3 | 0.277 | 0.24
| Std Dev. |  | 0.0092962 | 0.012394 | 0.004899 | 0.0064181 | 0.0075277 | 0 | 0.0028789 | 0.008165 | 0.022343 | 0
| Rel.Std.Dev. |  | 2.69% | 3.861% | 1.781% | 1.887% | 2.705% | 0% | 1.208% | 2.752% | 8.306% | 0%
| PDM3 |  | 21.26% | 12.65% | -3.492% | 19.39% | -2.322% | -15.77% | -16.37% | 4.112% | -5.597% | -15.77%
| Z-Score (Absolute) |  | 1.3 | 0.84 | 0.03 | 1.2 | 0.03 | 0.67 | 0.72 | 0.44 | 0.14 | 0.67
| Analytical results for La2O3 in OREAS 461 (Certified Value 3155 ppm) | [16 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 3166.6 | 2978.9 | 3178.3 | 3072.7 | 2838.2 | 2839.2725 | 3403.58 | 3116.1 | 2925.0658 | 2440.6 | 3080.9 | 3442.1863 | 3131.38 | 3249 | 3530 | 1
| 1 | 2 | 3154.8 | 3201.7 | 3107.9 | 2978.9 | 2896.8 | 2858.4295 | 3280.67 | 3127.9 | 2898.6554 | 2427.7 | 3123.2 | 3496.7239 | 3002.37 | 3213 | 3483 | #N/A
| 1 | 3 | 3248.7 | 3061 | 3166.6 | 3072.7 | 2990.6 | 2882.7398 | 3343.07 | 3109.1 | 2896.7039 | 2374.9 | 3160.7 | 3289.4471 | 3084.46 | 3237 | 3589 | 16
| 1 | 4 | 3178.3 | 3002.4 | 3236.9 | 3072.7 | 2978.9 | 2844.78 | 3289.59 | 3122 | 3021.4413 | 2356.2 | 3138.4 | 3410.5774 | 3190.02 | 3225 | 3518 | 3154.9337
| 1 | 5 | 3154.8 | 3225.2 | 3213.5 | 3096.2 | 2920.3 | 2913.7354 | 3295.1 | 3127.9 | 3000.3459 | 2386.6 | 3141.9 | 3346.6126 | 3201.74 | 3296 | 3636 | 36
| 1 | 6 | 3143.1 | 3061 | 3260.4 | 3119.6 | 2932 | 2900.1732 | 3246.31 | 3109.1 | 3012.5099 | 2338.6 | 3256.9 | 3393.0443 | 3166.56 | 3378 | 3483
| Mean |  | 3174.3833 | 3088.3667 | 3193.9333 | 3068.8 | 2926.1333 | 2873.1884 | 3309.72 | 3118.6833 | 2959.1204 | 2387.4333 | 3150.3333 | 3396.4319 | 3129.4217 | 3266.3333 | 3539.8333
| Median |  | 3160.7 | 3061 | 3195.9 | 3072.7 | 2926.15 | 2870.5847 | 3292.345 | 3119.05 | 2962.7058 | 2380.75 | 3140.15 | 3401.8109 | 3148.97 | 3243 | 3524
| Std Dev. |  | 38.3264 | 102.4099 | 54.8453 | 47.8728 | 55.8595 | 30.4586 | 55.5079 | 8.6168 | 58.5584 | 39.9151 | 58.7106 | 72.5138 | 75.4306 | 61.7468 | 61.1536
| Rel.Std.Dev. |  | 1.207% | 3.316% | 1.717% | 1.56% | 1.909% | 1.06% | 1.677% | 0.2763% | 1.979% | 1.672% | 1.864% | 2.135% | 2.41% | 1.89% | 1.728%
| PDM3 |  | 0.6165% | -2.11% | 1.236% | -2.73% | -7.252% | -8.93% | 4.906% | -1.149% | -6.207% | -24.33% | -0.1458% | 7.655% | -0.8086% | 3.531% | 12.2%
| Z-Score (Absolute) |  | 0.27 | 0.24 | 0.38 | 0.25 | 1.19 | 1.5 | 1.06 | 0.06 | 1 | 4.35 | 0 | 1.57 | 0 | 0.67 | 2.41
| Analytical results for Li in OREAS 461 (Certified Value 12 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 1
1
| 1 | 1 | 13.2694 | 13 | 12 | 10 | 12 | 10 | [9 empty] | 1
| 1 | 2 | 13.1907 | 13 | 13 | 10 | 13 | 10 | [9 empty] | #N/A
| 1 | 3 | 13.7971 | 13 | 12 | 10 | 12 | 10 | [9 empty] | 16
| 1 | 4 | 12.856 | 13 | 14 | < 10 | 13 | 10 | [9 empty] | 12.0055
| 1 | 5 | 14.3376 | 13 | 13 | 10 | 13 | 10 | [9 empty] | 37
| 1 | 6 | 14.7479 | 13 | 12 | 10 | 13 | 10
| Mean |  | 13.6998 | 13 | 12.6667 | 10 | 12.6667 | 10
| Median |  | 13.5332 | 13 | 12.5 | 10 | 13 | 10
| Std Dev. |  | 0.73091 | 0 | 0.8165 | 0 | 0.5164 | 0
| Rel.Std.Dev. |  | 5.335% | 0% | 6.446% | 0% | 4.077% | 0%
| PDM3 |  | 14.11% | 8.284% | 5.507% | -16.7% | 5.507% | -16.7%
| Z-Score (Absolute) |  | 1.02 | 0.33 | 0 | 2.63 | 0 | 2.63
| Analytical results for Lu2O3 in OREAS 461 (Certified Value 0.59 ppm) | [16 empty] | Certified
Lu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.591 | 0.569 | 0.557 | 0.591 | 0.557 | 0.59038 | 0.57 | 0.57 | 0.443 | 0.57 | < 1.1 | 0.341 | 0.68 | 0.785 |  | 1
| 1 | 2 | 0.591 | 0.716 | 0.569 | 0.58 | 0.546 | 0.5684 | 0.57 | 0.57 | 0.455 | 0.57 | < 1.1 | 0.33 | 0.68 | 0.785 |  | 1
| 1 | 3 | 0.637 | 0.625 | 0.512 | 0.591 | 0.523 | 0.63388 | 0.57 | 0.57 | 0.443 | 0.57 | < 1.1 | 0.33 | 0.68 | 0.716 |  | 16
| 1 | 4 | 0.614 | 0.614 | 0.569 | 0.625 | 0.557 | 0.57446 | 0.57 | 0.57 | 0.455 | 0.57 | < 1.1 | 0.478 | 0.68 | 0.944 |  | 0.58606
| 1 | 5 | 0.58 | 0.625 | 0.546 | 0.591 | 0.546 | 0.53728 | 0.68 | 0.45 | 0.455 | 0.57 | < 1.1 | 0.353 | 0.68 | 0.569 |  | 38
| 1 | 6 | 0.603 | 0.637 | 0.546 | 0.625 | 0.569 | 0.54546 | 0.45 | 0.57 | 0.432 | 0.57 | < 1.1 | 0.341 | 0.68 | 0.5
| Mean |  | 0.60267 | 0.631 | 0.54983 | 0.6005 | 0.54967 | 0.57498 | 0.56833 | 0.55 | 0.44717 | 0.57 |  | 0.36217 | 0.68 | 0.7165
| Median |  | 0.597 | 0.625 | 0.5515 | 0.591 | 0.5515 | 0.57143 | 0.57 | 0.57 | 0.449 | 0.57 |  | 0.341 | 0.68 | 0.7505
| Std Dev. |  | 0.020461 | 0.047887 | 0.021198 | 0.01945 | 0.015616 | 0.034763 | 0.072778 | 0.04899 | 0.0094745 | 0 |  | 0.057388 | 0 | 0.1611
| Rel.Std.Dev. |  | 3.395% | 7.589% | 3.855% | 3.239% | 2.841% | 6.046% | 12.81% | 8.907% | 2.119% | 0% |  | 15.85% | 0% | 22.48%
| PDM3 |  | 2.834% | 7.668% | -6.181% | 2.464% | -6.21% | -1.891% | -3.025% | -6.153% | -23.7% | -2.74% |  | -38.2% | 16.03% | 22.26%
| Z-Score (Absolute) |  | 0.71 | 0.95 | 0.4 | 0.67 | 0.4 | 0.13 | 0 | 0.4 | 2.54 | 0 | NA | 4.81 | 2.38 | 3.09
| Analytical results for MgO in OREAS 461 (Certified Value 1.74 wt.%) | [16 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.7414 | 1.89 | 1.774 | 1.7066 | 1.57 | 1.8323 | 1.651 | 1.68 | 1.791 | 1.807 | 1.774 |  |  |  |  | 1
| 1 | 2 | 1.7516 | 1.857 | 1.807 | 1.7269 | 1.61 | 1.8074 | 1.632 | 1.67 | 1.741 | 1.807 | 1.758 |  |  |  |  | 2
| 1 | 3 | 1.7499 | 1.857 | 1.741 | 1.7162 | 1.58 | 1.8074 | 1.652 | 1.69 | 1.658 | 1.774 | 1.691 |  |  |  |  | 16
| 1 | 4 | 1.7071 | 1.874 | 1.807 | 1.6926 | 1.6 | 1.8157 | 1.638 | 1.65 | 1.658 | 1.807 | 1.791 |  |  |  |  | 1.7361
| 1 | 5 | 1.7294 | 1.857 | 1.791 | 1.6965 | 1.57 | 1.7991 | 1.613 | 1.62 | 1.708 | 1.841 | 1.807 |  |  |  |  | 39
| 1 | 6 | 1.7208 | 1.824 | 1.741 | 1.7107 | 1.55 | 1.8157 | 1.649 | 1.68 | 1.774 | 1.857 | 1.791
| Mean |  | 1.7334 | 1.8598 | 1.7768 | 1.7083 | 1.58 | 1.8129 | 1.6392 | 1.665 | 1.7217 | 1.8155 | 1.7687
| Median |  | 1.7354 | 1.857 | 1.7825 | 1.7087 | 1.575 | 1.8115 | 1.6435 | 1.675 | 1.7245 | 1.807 | 1.7825
| Std Dev. |  | 0.017535 | 0.021995 | 0.030321 | 0.012676 | 0.021909 | 0.01134 | 0.015092 | 0.025884 | 0.056944 | 0.029365 | 0.041572
| Rel.Std.Dev. |  | 1.012% | 1.183% | 1.706% | 0.7421% | 1.387% | 0.6255% | 0.9207% | 1.555% | 3.308% | 1.617% | 2.35%
| PDM3 |  | -0.1595% | 7.125% | 2.344% | -1.606% | -8.993% | 4.424% | -5.585% | -4.097% | -0.8332% | 4.572% | 1.874%
| Z-Score (Absolute) |  | 0 | 1.25 | 0.43 | 0.25 | 1.51 | 0.79 | 0.93 | 0.67 | 0.11 | 0.81 | 0.5
| Analytical results for MnO in OREAS 461 (Certified Value 843 ppm) | [16 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [5 empty] | ppm
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | < 2600 | 900 | 897.4511 | 700 | 904 | 812 | 800 | 878 | 852 | 813 | [5 empty] | 1
| 1 | 2 | < 2600 | 900 | 916.1787 | 800 | 904 | 796 | 800 | 865 | 852 | 813 | [5 empty] | 3
| 1 | 3 | < 2600 | 900 | 911.3315 | 700 | 904 | 808 | 700 | 826 | 852 | 852 | [5 empty] | 16
| 1 | 4 | < 2600 | 900 | 904.9672 | 800 | 904 | 802 | 600 | 826 | 878 | 839 | [5 empty] | 842.6844
| 1 | 5 | < 2600 | 900 | 902.248 | 700 | 904 | 786 | 800 | 865 | 878 | 826 | [5 empty] | 40
| 1 | 6 | < 2600 | 900 | 904.2441 | 700 | 904 | 799 | 800 | 891 | 904 | 826
| Mean |  |  | 900 | 906.0701 | 733.3333 | 904 | 800.5 | 750 | 858.5 | 869.3333 | 828.1667
| Median |  |  | 900 | 904.6056 | 700 | 904 | 800.5 | 800 | 865 | 865 | 826
| Std Dev. |  |  | 0 | 6.683 | 51.6398 | 0 | 9.2033 | 83.666 | 26.9574 | 21.2289 | 15.1976
| Rel.Std.Dev. |  |  | 0% | 0.7376% | 7.042% | 0% | 1.15% | 11.16% | 3.14% | 2.442% | 1.835%
| PDM3 |  |  | 6.802% | 7.522% | -12.98% | 7.276% | -5.006% | -11% | 1.877% | 3.162% | -1.723%
| Z-Score (Absolute) |  | NA | 0.67 | 0.71 | 1.87 | 0.67 | 0.87 | 1.17 | 0 | 0.16 | 0.45
| Analytical results for Mo in OREAS 461 (Certified Value 45.7 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 1
1
| 1 | 1 | 47 | 24.8881 | 37 | 50 | 39.1087 | 46 | 50 | 48 | [7 empty] | 1
| 1 | 2 | 49 | 24.6069 | 37 | 50 | 41.6616 | 45 | 50 | 45 | [7 empty] | #N/A
| 1 | 3 | 47 | 23.9812 | 37 | 50 | 39.286 | 46 | 50 | 46 | [7 empty] | 16
| 1 | 4 | 47 | 24.9962 | 37 | 50 | 39.9942 | 46 | 50 | 49 | [7 empty] | 45.6795
| 1 | 5 | 48 | 24.4595 | 37 | 50 | 41.1223 | 47 | 50 | 54 | [7 empty] | 41
| 1 | 6 | 47 | 24.441 | 37 | 50 | 43.3661 | 46 | 50 | 49
| Mean |  | 47.5 | 24.5622 | 37 | 50 | 40.7565 | 46 | 50 | 48.5
| Median |  | 47 | 24.5332 | 37 | 50 | 40.5583 | 46 | 50 | 48.5
| Std Dev. |  | 0.83666 | 0.36313 | 0 | 0 | 1.6262 | 0.63246 | 0 | 3.1464
| Rel.Std.Dev. |  | 1.761% | 1.478% | 0% | 0% | 3.99% | 1.375% | 0% | 6.487%
| PDM3 |  | 3.985% | -46.23% | -19% | 9.458% | -10.78% | 0.7016% | 9.458% | 6.175%
| Z-Score (Absolute) |  | 0.16 | 4.6 | 2.02 | 0.67 | 1.24 | 0.16 | 0.67 | 0.36
| Analytical results for Na in OREAS 461 (Indicative Value 0.166 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 11 | 14 | 18 | 20 | [11 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES | BF*OES | [11 empty] | 3
3
| 1 | 1 | 0.17 | 0.171 | 0.15898 | 0.86 | [11 empty] | 1
| 1 | 2 | 0.17 | 0.178 | 0.16647 | 0.85 | [11 empty] | 16
| 1 | 3 | 0.17 | 0.171 | 0.15215 | 0.86 | [11 empty] | 16
| 1 | 4 | 0.17 | 0.178 | 0.15386 | 0.84 | [11 empty] | 0.16624
| 1 | 5 | 0.17 | 0.171 | 0.14555 | 0.77 | [11 empty] | 22
| 1 | 6 | 0.17 | 0.171 | 0.15675 | 0.83
| Mean |  | 0.17 | 0.17333 | 0.15563 | 0.835
| Median |  | 0.17 | 0.171 | 0.1553 | 0.845
| Std Dev. |  | 0 | 0.0036148 | 0.0070257 | 0.033912
| Rel.Std.Dev. |  | 0% | 2.085% | 4.515% | 4.061%
| PDM3 |  | 2.261% | 4.267% | -6.385% | 402.3%
| Z-Score (Absolute) |  | 0.12 | 0.12 | 1.23 | 52.22
| Analytical results for Nb2O5 in OREAS 461 (Certified Value 1854 ppm) | [16 empty] | Certified
Nb2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 1938.3 | 1952.6 | 1845.3 | 1938.3 | 1738.1 | 1959.8 | 1843.9 | > 1400 | 1587.9 | 1648.7221 | 1516.3 | 1874 | 1974.1 |  |  | 1
| 1 | 2 | 1931.2 | 2117.1 | 1859.6 | 1866.8 | 1759.5 | 1966.9 | 1862.5 | > 1400 | 1630.8 | 1660.4694 | 1429.1 | 1874 | 1931.2 |  |  | 5
| 1 | 3 | 1952.6 | 2002.7 | 1831 | 1931.2 | 1823.9 | 2011.3 | 1833.9 | > 1400 | 1587.9 | 1486.8143 | 1379 | 1874 | 2074.2 |  |  | 16
| 1 | 4 | 1909.7 | 1938.3 | 1902.6 | 1938.3 | 1816.7 | 1984.1 | 1848.2 | > 1400 | 1630.8 | 1639.6394 | 1253.1 | 1802.4 | 1974.1 |  |  | 1854.4844
| 1 | 5 | 1909.7 | 2059.9 | 1888.3 | 1931.2 | 1781 | 1958.4 | 1853.9 | > 1400 | 1723.8 | 1580.7812 | 1530.6 | 1888.3 | 1988.4 |  |  | 42
| 1 | 6 | 1931.2 | 1924 | 1916.9 | 1959.8 | 1766.7 | 1959.8 | 1818.2 | -- | 1516.3 | 1580.3263 | 1502 | 1945.5 | 1988.4
| Mean |  | 1928.7833 | 1999.1 | 1873.95 | 1927.6 | 1780.9833 | 1973.3833 | 1843.4333 |  | 1612.9167 | 1599.4588 | 1435.0167 | 1876.3667 | 1988.4
| Median |  | 1931.2 | 1977.65 | 1873.95 | 1934.75 | 1773.85 | 1963.35 | 1846.05 |  | 1609.35 | 1610.2103 | 1465.55 | 1874 | 1981.25
| Std Dev. |  | 16.7202 | 76.3315 | 33.8907 | 31.586 | 33.5322 | 20.9192 | 15.6437 |  | 68.5874 | 65.0651 | 106.2981 | 45.6283 | 46.9946
| Rel.Std.Dev. |  | 0.8669% | 3.818% | 1.809% | 1.639% | 1.883% | 1.06% | 0.8486% |  | 4.252% | 4.068% | 7.407% | 2.432% | 2.363%
| PDM3 |  | 4.006% | 7.798% | 1.05% | 3.943% | -3.963% | 6.411% | -0.5959% |  | -13.03% | -13.75% | -22.62% | 1.18% | 7.221%
| Z-Score (Absolute) |  | 0.38 | 0.88 | 0.01 | 0.46 | 0.67 | 0.7 | 0.22 | NA | 1.86 | 1.95 | 3.12 | 0.01 | 0.68
| Analytical results for Nd2O3 in OREAS 461 (Certified Value 1900 ppm) | [16 empty] | Certified
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1691.3 | 1889.6 | 1942.1 | 1930.4 | 1755.4 | 1880.9984 | 2013.2 | 1877.9 | 1587.216 | 1597 | 1895.4 | 1807.2549 | 1924.56 | 1971.2 | 2122.8 | 1
| 1 | 2 | 1702.9 | 2023.7 | 1918.7 | 1866.2 | 1807.9 | 1873.9345 | 1992.2 | 1854.23 | 1647.7 | 1592 | 1922.2 | 1815.4227 | 1842.91 | 1772.9 | 2122.8 | #N/A
| 1 | 3 | 1778.8 | 1895.4 | 1953.7 | 1912.9 | 1854.6 | 1893.034 | 1989.9 | 1846.53 | 1590.4977 | 1563 | 1928.1 | 1779.5905 | 1912.9 | 1819.6 | 2297.8 | 16
| 1 | 4 | 1732.1 | 1895.4 | 1936.2 | 1918.7 | 1842.9 | 1874.4065 | 1980.5 | 1897.27 | 1603.4197 | 1549 | 1929.2 | 1807.0104 | 1971.22 | 1807.9 | 2111.2 | 1899.8093
| 1 | 5 | 1691.3 | 2023.7 | 1959.6 | 1924.6 | 1860.4 | 1934.6737 | 1993.4 | 1855.86 | 1634.6222 | 1589 | 1936.2 | 1792.575 | 1971.22 | 1854.6 | 2134.5 | 43
| 1 | 6 | 1679.6 | 1936.2 | 1912.9 | 1947.9 | 1866.2 | 1911.7549 | 1946.7 | 1828.22 | 1566.0775 | 1550 | 1974.7 | 1778.8136 | 1959.55 | 1901.2 | 2157.8
| Mean |  | 1712.6667 | 1944 | 1937.2 | 1916.7833 | 1831.2333 | 1894.8003 | 1985.9833 | 1860.0017 | 1604.9222 | 1573.3333 | 1930.9667 | 1796.7779 | 1930.3933 | 1854.5667 | 2157.8167
| Median |  | 1697.1 | 1915.8 | 1939.15 | 1921.65 | 1848.75 | 1887.0162 | 1991.05 | 1855.045 | 1596.9587 | 1576 | 1928.65 | 1799.7927 | 1942.055 | 1837.1 | 2128.65
| Std Dev. |  | 37.022 | 63.9504 | 18.6105 | 27.5397 | 42.5455 | 24.1754 | 22.0136 | 24.2953 | 30.8051 | 21.8967 | 25.6639 | 15.4783 | 49.35 | 71.8998 | 70.3768
| Rel.Std.Dev. |  | 2.162% | 3.29% | 0.9607% | 1.437% | 2.323% | 1.276% | 1.108% | 1.306% | 1.919% | 1.392% | 1.329% | 0.8614% | 2.556% | 3.877% | 3.261%
| PDM3 |  | -9.851% | 2.326% | 1.968% | 0.8935% | -3.61% | -0.2637% | 4.536% | -2.095% | -15.52% | -17.18% | 1.64% | -5.423% | 1.61% | -2.381% | 13.58%
| Z-Score (Absolute) |  | 2.72 | 0.69 | 0.59 | 0.31 | 0.67 | 0 | 1.27 | 0.48 | 4.04 | 4.48 | 0.5 | 1.37 | 0.5 | 0.56 | 3.28
| Analytical results for NiO in OREAS 461 (Certified Value 81 ppm) | [16 empty] | Certified
NiO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
1
| 1 | 1 | 87.8 | 80.2 | < 130 | 86.1744 | 78.9 | 89 | 127 | 78.9 | [7 empty] | 1
| 1 | 2 | 85.3 | 82.7 | < 130 | 91.6453 | 76.4 | 76 | 127 | 84 | [7 empty] | #N/A
| 1 | 3 | 80.2 | 78.9 | < 130 | 80.8321 | 71.3 | 76 | 127 | 94.2 | [7 empty] | 16
| 1 | 4 | 78.9 | 87.8 | < 130 | 80.7198 | 66.2 | 76 | 127 | 89.1 | [7 empty] | 81.0513
| 1 | 5 | 91.6 | 78.9 | < 130 | 77.6075 | 78.9 | 76 | 127 | 77.6 | [7 empty] | 44
| 1 | 6 | 61.1 | 81.4 | < 130 | 89.4791 | 78.9 | 89 | 127 | 78.9
| Mean |  | 80.8167 | 81.65 |  | 84.4097 | 75.1 | 80.3333 | 127 | 83.7833
| Median |  | 82.75 | 80.8 |  | 83.5033 | 77.65 | 76 | 127 | 81.45
| Std Dev. |  | 10.7513 | 3.3531 |  | 5.5471 | 5.2661 | 6.7132 | 0 | 6.6746
| Rel.Std.Dev. |  | 13.3% | 4.107% |  | 6.572% | 7.012% | 8.357% | 0% | 7.966%
| PDM3 |  | -0.2895% | 0.7387% |  | 4.144% | -7.343% | -0.8858% | 56.69% | 3.371%
| Z-Score (Absolute) |  | 0.27 | 0 | NA | 0.88 | 2.09 | 0.34 | 14.5 | 0.67
| Analytical results for P2O5 in OREAS 461 (Certified Value 0.822 wt.%) | [16 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.8754 | 0.871 | 0.848 | 0.88323 | 0.6 | 0.9625 | 0.8324 | 0.756 | 0.779 | 0.756 | [5 empty] | 1
| 1 | 2 | 0.8731 | 0.871 | 0.848 | 0.87157 | 0.61 | 0.7906 | 0.8351 | 0.733 | 0.779 | 0.756 | [5 empty] | 6
| 1 | 3 | 0.8525 | 0.848 | 0.871 | 0.87354 | 0.6 | 0.8135 | 0.8415 | 0.779 | 0.779 | 0.779 | [5 empty] | 16
| 1 | 4 | 0.8319 | 0.871 | 0.825 | 0.88969 | 0.63 | 0.8021 | 0.8204 | 0.756 | 0.779 | 0.756 | [5 empty] | 0.82204
| 1 | 5 | 0.8708 | 0.848 | 0.848 | 0.84801 | 0.62 | 0.7906 | 0.8502 | 0.756 | 0.825 | 0.779 | [5 empty] | 45
| 1 | 6 | 0.8731 | 0.848 | 0.825 | 0.85223 | 0.57 | 0.8364 | 0.8078 | 0.756 | 0.825 | 0.779
| Mean |  | 0.8628 | 0.8595 | 0.84417 | 0.86971 | 0.605 | 0.83262 | 0.83123 | 0.756 | 0.79433 | 0.7675
| Median |  | 0.87195 | 0.8595 | 0.848 | 0.87255 | 0.605 | 0.8078 | 0.83375 | 0.756 | 0.779 | 0.7675
| Std Dev. |  | 0.017296 | 0.012598 | 0.017314 | 0.016595 | 0.020736 | 0.065896 | 0.015152 | 0.014546 | 0.023754 | 0.012598
| Rel.Std.Dev. |  | 2.005% | 1.466% | 2.051% | 1.908% | 3.428% | 7.914% | 1.823% | 1.924% | 2.99% | 1.641%
| PDM3 |  | 4.959% | 4.558% | 2.692% | 5.8% | -26.4% | 1.287% | 1.119% | -8.033% | -3.37% | -6.634%
| Z-Score (Absolute) |  | 0.75 | 0.6 | 0.37 | 0.76 | 3.19 | 0.18 | 0.18 | 0.94 | 0.37 | 0.76
| Analytical results for PbO in OREAS 461 (Certified Value 114 ppm) | [16 empty] | Certified
PbO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 113.1 | 112 | 110 | 114.6776 | 61.4 | < 320 | 118 | 107.7 | [7 empty] | 1
| 1 | 2 | 114.2 | 107.7 | 220 | 114.9415 | 54.9 | < 320 | 118 | 109.9 | [7 empty] | #N/A
| 1 | 3 | 111 | 107.7 | 220 | 111.4587 | 54.9 | < 320 | 129 | 115.3 | [7 empty] | 16
| 1 | 4 | 116.3 | 99.1 | < 110 | 113.4107 | 62.5 | < 320 | 118 | 119.6 | [7 empty] | 114.4986
| 1 | 5 | 113.1 | 106.6 | 110 | 109.708 | 54.9 | < 320 | 129 | 134.6 | [7 empty] | 46
| 1 | 6 | 112 | 120.6 | 110 | 112.6983 | 57.1 | < 320 | 129 | 115.3
| Mean |  | 113.2833 | 108.95 | 154 | 112.8158 | 57.6167 |  | 123.5 | 117.0667
| Median |  | 113.1 | 107.7 | 110 | 113.0545 | 56 |  | 123.5 | 115.3
| Std Dev. |  | 1.8368 | 7.0809 | 60.2495 | 1.9926 | 3.4805 |  | 6.0249 | 9.583
| Rel.Std.Dev. |  | 1.621% | 6.499% | 39.12% | 1.766% | 6.041% |  | 4.879% | 8.186%
| PDM3 |  | -1.061% | -4.846% | 34.5% | -1.47% | -49.68% |  | 7.862% | 2.243%
| Z-Score (Absolute) |  | 0 | 0.67 | 5.87 | 0.07 | 8.71 | NA | 1.66 | 0.04
| Analytical results for Pr6O11 in OREAS 461 (Certified Value 590 ppm) | [16 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 530.4 | 589.6 | 589.6 | 586 | 554.6 | 603.0422 | 601.7 | 618.72 | 438.3468 | 464.601 | 574.5 | 601.3726 | 582.352 | 676.6 | 573.9 | 1
| 1 | 2 | 534 | 631.9 | 575.1 | 569.1 | 579.9 | 618.4089 | 544.9 | 624.52 | 441.4356 | 458.258 | 578.73 | 603.1428 | 554.564 | 605.3 | 569.1 | #N/A
| 1 | 3 | 549.7 | 598.1 | 596.9 | 584.8 | 594.4 | 635.5957 | 572.7 | 614.85 | 436.0532 | 448.496 | 584.16 | 576.4031 | 571.479 | 617.4 | 670.6 | 16
| 1 | 4 | 537.6 | 595.6 | 593.2 | 592 | 586 | 618.5843 | 559.4 | 630.32 | 440.2627 | 445.415 | 586.58 | 587.6758 | 590.81 | 608.9 | 566.6 | 590.314
| 1 | 5 | 541.3 | 612.6 | 601.7 | 586 | 594.4 | 639.6767 | 616.2 | 607.12 | 436.4134 | 460.687 | 585.37 | 580.7489 | 588.393 | 623.4 | 575.1 | 47
| 1 | 6 | 535.2 | 587.2 | 602.9 | 596.9 | 592 | 633.7451 | 575.1 | 609.05 | 436.6615 | 442.6 | 601.68 | 587.5222 | 584.769 | 653.6 | 592
| Mean |  | 538.0333 | 602.5 | 593.2333 | 585.8 | 583.55 | 624.8421 | 578.3333 | 617.43 | 438.1955 | 453.3428 | 585.17 | 589.4776 | 578.7278 | 630.8667 | 591.2167
| Median |  | 536.4 | 596.85 | 595.05 | 586 | 589 | 626.1647 | 573.9 | 616.785 | 437.5041 | 453.377 | 584.765 | 587.599 | 583.5605 | 620.4 | 574.5
| Std Dev. |  | 6.7754 | 16.9342 | 10.2039 | 9.3981 | 15.2596 | 13.9338 | 26.4512 | 8.9584 | 2.2322 | 9.0177 | 9.2839 | 10.7929 | 13.6049 | 28.2193 | 39.8925
| Rel.Std.Dev. |  | 1.259% | 2.811% | 1.72% | 1.604% | 2.615% | 2.23% | 4.574% | 1.451% | 0.5094% | 1.989% | 1.587% | 1.831% | 2.351% | 4.473% | 6.748%
| PDM3 |  | -8.856% | 2.064% | 0.4945% | -0.7647% | -1.146% | 5.849% | -2.03% | 4.593% | -25.77% | -23.2% | -0.8714% | -0.1417% | -1.963% | 6.87% | 0.1529%
| Z-Score (Absolute) |  | 3.09 | 0.7 | 0.48 | 0 | 0.23 | 2.53 | 0.48 | 2.05 | 9.54 | 8.56 | 0.04 | 0.24 | 0.46 | 2.32 | 0.67
| Analytical results for Rb in OREAS 461 (Certified Value 13.5 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  | 1
2
| 1 | 1 | 13.4 | 13.4 | 13.7 | 12.9 | 12.4 | 13.5921 | 13.7 | 14.5 | 13.8667 | 11.2 | 13 | 11.6 | 19 | 13.2 |  | 1
| 1 | 2 | 12.9 | 14.9 | 14 | 12.6 | 13.2 | 13.6904 | 13.5 | 14.2 | 13.7634 | 10.4 | 14 | 11.2 | 18 | 13 |  | #N/A
| 1 | 3 | 13.7 | 13.6 | 14 | 13.2 | 12.7 | 13.5146 | 13.7 | 14.2 | 13.8084 | 10.7 | 13 | 11.3 | 16.5 | 13.1 |  | 16
| 1 | 4 | 12.5 | 13.4 | 13.9 | 13.1 | 13.4 | 13.456 | 14 | 14.3 | 13.8893 | 10.3 | 14 | 11.9 | 15.5 | 13.5 |  | 13.5154
| 1 | 5 | 12.2 | 13.9 | 14.3 | 13.3 | 13.6 | 13.1719 | 13.1 | 14.2 | 13.6644 | 10.1 | 15 | 11.7 | 15.5 | 14.2 |  | 48
| 1 | 6 | 12.7 | 13.5 | 14 | 13.1 | 12.9 | 13.7205 | 13.7 | 13.9 | 14.0615 | 9.9 | 18 | 11.3 | 16.5 | 13
| Mean |  | 12.9 | 13.7833 | 13.9833 | 13.0333 | 13.0333 | 13.5242 | 13.6167 | 14.2167 | 13.8423 | 10.4333 | 14.5 | 11.5 | 16.8333 | 13.3333
| Median |  | 12.8 | 13.55 | 14 | 13.1 | 13.05 | 13.5534 | 13.7 | 14.2 | 13.8376 | 10.35 | 14 | 11.45 | 16.5 | 13.15
| Std Dev. |  | 0.56214 | 0.57764 | 0.19408 | 0.25033 | 0.45019 | 0.19982 | 0.29944 | 0.19408 | 0.13408 | 0.46332 | 1.8708 | 0.27568 | 1.4024 | 0.46332
| Rel.Std.Dev. |  | 4.358% | 4.191% | 1.388% | 1.921% | 3.454% | 1.477% | 2.199% | 1.365% | 0.9687% | 4.441% | 12.9% | 2.397% | 8.331% | 3.475%
| PDM3 |  | -4.554% | 1.982% | 3.462% | -3.567% | -3.567% | 0.06515% | 0.749% | 5.188% | 2.418% | -22.8% | 7.285% | -14.91% | 24.55% | -1.347%
| Z-Score (Absolute) |  | 0.91 | 0.03 | 0.63 | 0.72 | 0.72 | 0.03 | 0.11 | 0.96 | 0.43 | 4.41 | 0.37 | 2.9 | 4.67 | 0.54
| Analytical results for Re in OREAS 461 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 17 | 22 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | [12 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 1 | [12 empty] | 17
| 1 | 3 | < 0.1 | < 0.1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | [12 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | [12 empty] | 23
| 1 | 6 | < 0.1 | < 0.1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 461 (Certified Value 518 ppm) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 0
0
| 1 | 1 | 410 | 500 | 400 | 700 | 600 | [10 empty] | 1
| 1 | 2 | 450 | 600 | 400 | 600 | 600 | [10 empty] | #N/A
| 1 | 3 | 420 | 500 | 500 | 600 | 400 | [10 empty] | 16
| 1 | 4 | 410 | 500 | 500 | 600 | 600 | [10 empty] | 517.6667
| 1 | 5 | 460 | 500 | 500 | 500 | 600 | [10 empty] | 49
| 1 | 6 | 460 | < 500 | 500 | 400 | 600
| Mean |  | 435 | 520 | 466.6667 | 566.6667 | 566.6667
| Median |  | 435 | 500 | 500 | 600 | 600
| Std Dev. |  | 24.2899 | 44.7214 | 51.6398 | 103.2796 | 81.6497
| Rel.Std.Dev. |  | 5.584% | 8.6% | 11.07% | 18.23% | 14.41%
| PDM3 |  | -15.97% | 0.4507% | -9.852% | 9.466% | 9.466%
| Z-Score (Absolute) |  | 1.07 | 0 | 0.67 | 0.59 | 1.01
| Analytical results for Sb in OREAS 461 (Certified Value 2.76 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 2.7 | 2.9 | 3 | 2.6073 | 3 | 2.6 | [9 empty] | 1
| 1 | 2 | 2.6 | 2.8 | 3 | 2.5343 | 3 | 2.4 | [9 empty] | #N/A
| 1 | 3 | 2.6 | 2.8 | 3 | 2.3028 | 3 | 2.7 | [9 empty] | 16
| 1 | 4 | 2.7 | 2.7 | 3 | 2.3942 | 3 | 2.8 | [9 empty] | 2.7613
| 1 | 5 | 2.7 | 2.8 | 3 | 2.4103 | 3 | 2.7 | [9 empty] | 50
| 1 | 6 | 2.6 | 2.8 | 3 | 2.3593 | 3 | 2.9
| Mean |  | 2.65 | 2.8 | 3 | 2.4347 | 3 | 2.6833
| Median |  | 2.65 | 2.8 | 3 | 2.4023 | 3 | 2.7
| Std Dev. |  | 0.054772 | 0.063246 | 0 | 0.11405 | 0 | 0.17224
| Rel.Std.Dev. |  | 2.067% | 2.259% | 0% | 4.684% | 0% | 6.419%
| PDM3 |  | -4.032% | 1.4% | 8.643% | -11.83% | 8.643% | -2.825%
| Z-Score (Absolute) |  | 0.35 | 0.22 | 1 | 1.18 | 1 | 0.22
| Analytical results for Sc in OREAS 461 (Indicative Value 44 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | [11 empty] | 1
1
| 1 | 1 | 42 | 50 | 40 | 110 | [11 empty] | 1
| 1 | 2 | 43 | 50 | 40 | 100 | [11 empty] | 18
| 1 | 3 | 42 | 50 | 40 | 120 | [11 empty] | 16
| 1 | 4 | 42 | 50 | 40 | 110 | [11 empty] | 44
| 1 | 5 | 42 | 50 | 40 | 110 | [11 empty] | 24
| 1 | 6 | 41 | 50 | 40 | 110
| Mean |  | 42 | 50 | 40 | 110
| Median |  | 42 | 50 | 40 | 110
| Std Dev. |  | 0.63246 | 0 | 0 | 6.3246
| Rel.Std.Dev. |  | 1.506% | 0% | 0% | 5.75%
| PDM3 |  | -4.545% | 13.64% | -9.091% | 150%
| Z-Score (Absolute) |  | 0.54 | 0.54 | 0.81 | 8.36
| Analytical results for Se in OREAS 461 (Indicative Value < 20 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 1
1
| 1 | 1 | < 20 | < 20 | [13 empty] | 1
| 1 | 2 | < 20 | < 20 | [13 empty] | 15
| 1 | 3 | < 20 | < 20 | [13 empty] | 16
| 1 | 4 | < 20 | < 20 | [13 empty] | < 20
| 1 | 5 | < 20 | < 20 | [13 empty] | 21
| 1 | 6 | < 20 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for SiO2 in OREAS 461 (Certified Value 31.46 wt.%) | [16 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 32.0878 | 30.826 | 33.0095 | 30.48 | 31.874 | 31.0193 | 28.5 | 30.804 | [7 empty] | 1
| 1 | 2 | 32.5532 | 30.911 | 32.5672 | 31.26 | 31.553 | 30.9123 | 28.6 | 30.804 | [7 empty] | 7
| 1 | 3 | 32.4262 | 30.633 | 31.9382 | 30.55 | 31.831 | 31.0835 | 28.1 | 30.163 | [7 empty] | 16
| 1 | 4 | 32.7647 | 30.911 | 32.1649 | 31.77 | 31.789 | 31.3403 | 28.1 | 30.804 | [7 empty] | 31.4637
| 1 | 5 | 32.7435 | 31.04 | 32.2078 | 30.11 | 31.553 | 31.0407 | 27.7 | 31.446 | [7 empty] | 51
| 1 | 6 | 32.0666 | 30.869 | 32.2665 | 30.05 | 31.81 | 31.565 | 28.2 | 31.874
| Mean |  | 32.4403 | 30.865 | 32.359 | 30.7033 | 31.735 | 31.1602 | 28.2 | 30.9825
| Median |  | 32.4897 | 30.89 | 32.2372 | 30.515 | 31.7995 | 31.0621 | 28.15 | 30.804
| Std Dev. |  | 0.30791 | 0.13433 | 0.37753 | 0.67822 | 0.14375 | 0.24404 | 0.32249 | 0.59612
| Rel.Std.Dev. |  | 0.9492% | 0.4352% | 1.167% | 2.209% | 0.453% | 0.7832% | 1.144% | 1.924%
| PDM3 |  | 3.104% | -1.903% | 2.846% | -2.417% | 0.8623% | -0.9646% | -10.37% | -1.529%
| Z-Score (Absolute) |  | 1.79 | 0.27 | 1.68 | 0.48 | 0.87 | 0.12 | 3.75 | 0.12
| Analytical results for Sm2O3 in OREAS 461 (Certified Value 255 ppm) | [16 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 248.15 | 257.43 | 251.63 | 252.79 | 229.6 | 254.8789 | 263 | 256.27 | 199.6857 | 205.94 | 243.52 | 259.5822 | 260.91 | 278.3 | 259.8 | 1
| 1 | 2 | 245.84 | 275.98 | 250.47 | 246.99 | 231.92 | 248.7704 | 256.97 | 257.32 | 203.7682 | 204.44 | 245.26 | 263.3283 | 246.99 | 254 | 252.8 | 8
| 1 | 3 | 255.11 | 256.27 | 256.27 | 253.95 | 241.2 | 250.0011 | 262.07 | 254.76 | 199.4638 | 201.42 | 246.42 | 250.4824 | 253.95 | 251.6 | 287.6 | 16
| 1 | 4 | 251.63 | 256.27 | 255.11 | 256.27 | 238.88 | 250.799 | 256.39 | 259.29 | 202.9513 | 199.45 | 246.99 | 259.0127 | 265.55 | 251.6 | 259.8 | 254.7024
| 1 | 5 | 249.31 | 275.98 | 259.75 | 255.11 | 242.36 | 250.99 | 255.46 | 253.72 | 205.0469 | 203.74 | 251.63 | 254.3966 | 266.71 | 262.1 | 257.4 | 52
| 1 | 6 | 246.99 | 257.43 | 252.79 | 260.91 | 242.36 | 254.4653 | 253.95 | 254.18 | 198.5069 | 200.49 | 256.85 | 257.3917 | 262.07 | 270.2 | 270.2
| Mean |  | 249.505 | 263.2267 | 254.3367 | 254.3367 | 237.72 | 251.6508 | 257.9733 | 255.9233 | 201.5704 | 202.58 | 248.445 | 257.3657 | 259.3633 | 261.3 | 264.6
| Median |  | 248.73 | 257.43 | 253.95 | 254.53 | 240.04 | 250.8945 | 256.68 | 255.515 | 201.3185 | 202.58 | 246.705 | 258.2022 | 261.49 | 258.05 | 259.8
| Std Dev. |  | 3.393 | 9.8923 | 3.4149 | 4.562 | 5.5873 | 2.4707 | 3.6897 | 2.1285 | 2.6907 | 2.5142 | 4.9276 | 4.4581 | 7.5415 | 11.0483 | 12.6295
| Rel.Std.Dev. |  | 1.36% | 3.758% | 1.343% | 1.794% | 2.35% | 0.9818% | 1.43% | 0.8317% | 1.335% | 1.241% | 1.983% | 1.732% | 2.908% | 4.228% | 4.773%
| PDM3 |  | -2.041% | 3.347% | -0.1436% | -0.1436% | -6.668% | -1.198% | 1.284% | 0.4793% | -20.86% | -20.46% | -2.457% | 1.046% | 1.83% | 2.59% | 3.886%
| Z-Score (Absolute) |  | 0.65 | 1.19 | 0 | 0 | 2.23 | 0.36 | 0.49 | 0.21 | 7.08 | 6.95 | 0.79 | 0.41 | 0.67 | 0.93 | 0.76
| Analytical results for SnO2 in OREAS 461 (Certified Value 32.5 ppm) | [16 empty] | Certified
SnO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS |  | 1
1
| 1 | 1 | 38.1 | 33 | 36.8 | 34.3 | 33 | 34.3 | 27.9 | 26.66 | 29.2 | 22.5251 | 130.8 | < 63 | 51 | 30.5 |  | 1
| 1 | 2 | 39.4 | 35.5 | 35.5 | 33 | 33 | 36.8 | 26.7 | 26.53 | 30.5 | 22.2041 | 33 | < 63 | 51 | 29.2 |  | 9
| 1 | 3 | 40.6 | 33 | 35.5 | 34.3 | 34.3 | 35.5 | 26.7 | 24.63 | 31.7 | 21.3308 | 31.7 | < 63 | 51 | 35.5 |  | 16
| 1 | 4 | 38.1 | 31.7 | 38.1 | 34.3 | 35.5 | 35.5 | 26.7 | 23.74 | 30.5 | 21.4266 | 30.5 | < 63 | 38 | 27.9 |  | 32.5402
| 1 | 5 | 38.1 | 34.3 | 35.5 | 35.5 | 33 | 35.5 | 27.9 | 23.87 | 30.5 | 23.6828 | 35.5 | < 63 | 51 | 27.9 |  | 53
| 1 | 6 | 39.4 | 33 | 36.8 | 36.8 | 34.3 | 36.8 | 26.7 | 23.87 | 31.7 | 20.3046 | 35.5 | < 63 | 51 | 29.2
| Mean |  | 38.95 | 33.4167 | 36.3667 | 34.7 | 33.85 | 35.7333 | 27.1 | 24.8833 | 30.6833 | 21.9123 | 49.5 |  | 48.8333 | 30.0333
| Median |  | 38.75 | 33 | 36.15 | 34.3 | 33.65 | 35.5 | 26.7 | 24.25 | 30.5 | 21.8154 | 34.25 |  | 51 | 29.2
| Std Dev. |  | 1.0291 | 1.3106 | 1.0614 | 1.2977 | 1.0291 | 0.94798 | 0.61968 | 1.3633 | 0.9347 | 1.1623 | 39.8793 |  | 5.3072 | 2.8493
| Rel.Std.Dev. |  | 2.642% | 3.922% | 2.919% | 3.74% | 3.04% | 2.653% | 2.287% | 5.479% | 3.046% | 5.304% | 80.56% |  | 10.87% | 9.487%
| PDM3 |  | 19.7% | 2.693% | 11.76% | 6.637% | 4.025% | 9.813% | -16.72% | -23.53% | -5.706% | -32.66% | 52.12% |  | 50.07% | -7.704%
| Z-Score (Absolute) |  | 1.25 | 0 | 0.67 | 0.29 | 0.1 | 0.53 | 1.44 | 1.95 | 0.63 | 2.62 | 0.04 | NA | 3.47 | 1.02
| Analytical results for SrO in OREAS 461 (Certified Value 685 ppm) | [16 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 682.4 | 659.9 | 675.3 | 678.8 | 669.4 | 711.9 | 678.1 | 645.7 | 710 | 701.252 | 630.33 | 689.5 | 674 |  |  | 1
| 1 | 2 | 678.8 | 709.6 | 684.7 | 655.2 | 670.5 | 719 | 676.33 | 675.3 | 710 | 734.8585 | 594.85 | 681.2 | 674 |  |  | 11
| 1 | 3 | 710.7 | 678.8 | 697.7 | 682.4 | 680 | 713.1 | 670.42 | 654 | 710 | 714.5452 | 614.95 | 700.1 | 639 |  |  | 16
| 1 | 4 | 678.8 | 662.3 | 685.9 | 678.8 | 674.1 | 704.8 | 686.38 | 678.8 | 710 | 739.0849 | 629.14 | 687.1 | 674 |  |  | 684.6537
| 1 | 5 | 680 | 697.7 | 706 | 683.5 | 691.8 | 714.3 | 673.14 | 638.6 | 710 | 769.7709 | 632.69 | 701.3 | 686 |  |  | 54
| 1 | 6 | 680 | 667 | 684.7 | 683.5 | 682.4 | 695.4 | 670.53 | 639.8 | 710 | 796.9227 | 618.5 | 706 | 686
| Mean |  | 685.1167 | 679.2167 | 689.05 | 677.0333 | 678.0333 | 709.75 | 675.8167 | 655.3667 | 710 | 742.739 | 620.0767 | 694.2 | 672.1667
| Median |  | 680 | 672.9 | 685.3 | 680.6 | 677.05 | 712.5 | 674.735 | 649.85 | 710 | 736.9717 | 623.82 | 694.8 | 674
| Std Dev. |  | 12.602 | 20.3642 | 10.9449 | 10.9121 | 8.476 | 8.3968 | 6.0216 | 17.693 | 0 | 35.3783 | 14.2234 | 9.6536 | 17.2791
| Rel.Std.Dev. |  | 1.839% | 2.998% | 1.588% | 1.612% | 1.25% | 1.183% | 0.891% | 2.7% | 0% | 4.763% | 2.294% | 1.391% | 2.571%
| PDM3 |  | 0.06762% | -0.7941% | 0.6421% | -1.113% | -0.967% | 3.666% | -1.291% | -4.278% | 3.702% | 8.484% | -9.432% | 1.394% | -1.824%
| Z-Score (Absolute) |  | 0 | 0.06 | 0.67 | 0.11 | 0.15 | 2.21 | 0.31 | 1.83 | 2.2 | 4.67 | 4.46 | 1.06 | 0.09
| Analytical results for Ta2O5 in OREAS 461 (Certified Value 30.7 ppm) | [16 empty] | Certified
Ta2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 1
2
| 1 | 1 | 28.57 | 29.18 | 28.94 | 29.55 | 31.02 | 30.1064 | 29.92 | 22.505 | 25.6 | 29.0101 | 34.68 | 31.14 | 30.53 |  |  | 1
| 1 | 2 | 28.94 | 32.11 | 30.28 | 29.92 | 31.14 | 30.4057 | 30.16 | 22.285 | 26.9 | 25.8934 | 32.6 | 31.14 | 31.63 |  |  | 1
| 1 | 3 | 29.92 | 31.38 | 29.79 | 30.77 | 31.87 | 30.5984 | 30.89 | 21.65 | 20.8 | 24.2558 | 31.5 | 30.53 | 37.24 |  |  | 16
| 1 | 4 | 29.18 | 28.21 | 31.02 | 30.65 | 32.11 | 30.6732 | 30.77 | 21.479 | 28.1 | 25.8941 | 29.06 | 31.14 | 30.77 |  |  | 30.6885
| 1 | 5 | 28.94 | 32.11 | 30.77 | 30.77 | 31.99 | 29.753 | 30.53 | 21.882 | 26.9 | 22.0239 | 34.92 | 32.36 | 30.65 |  |  | 55
| 1 | 6 | 28.82 | 29.31 | 29.43 | 30.89 | 31.75 | 29.9788 | 29.92 | 21.589 | 20.8 | 24.0512 | 34.31 | 32.97 | 31.5
| Mean |  | 29.0617 | 30.3833 | 30.0383 | 30.425 | 31.6467 | 30.2526 | 30.365 | 21.8983 | 24.85 | 25.1881 | 32.845 | 31.5467 | 32.0533
| Median |  | 28.94 | 30.345 | 30.035 | 30.71 | 31.81 | 30.2561 | 30.345 | 21.766 | 26.25 | 25.0746 | 33.455 | 31.14 | 31.135
| Std Dev. |  | 0.46486 | 1.6899 | 0.79944 | 0.55237 | 0.45663 | 0.36492 | 0.42543 | 0.41256 | 3.2353 | 2.357 | 2.281 | 0.91838 | 2.5816
| Rel.Std.Dev. |  | 1.6% | 5.562% | 2.661% | 1.816% | 1.443% | 1.206% | 1.401% | 1.884% | 13.02% | 9.358% | 6.945% | 2.911% | 8.054%
| PDM3 |  | -5.301% | -0.9943% | -2.118% | -0.8585% | 3.122% | -1.42% | -1.054% | -28.64% | -19.02% | -17.92% | 7.027% | 2.796% | 4.447%
| Z-Score (Absolute) |  | 0.98 | 0.02 | 0.25 | 0.05 | 0.97 | 0.08 | 0 | 6.38 | 4.17 | 3.9 | 1.56 | 0.67 | 0.49
| Analytical results for Tb4O7 in OREAS 461 (Certified Value 10.7 ppm) | [16 empty] | Certified
Tb4O7 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 10.151 | 10.586 | 9.962 | 10.209 | 10.315 | 11.0845 | 11.06 | 10.82 | 8.2135 | 9.845 | 10.59 | 13.8349 | 10.68 | 12.94 | 10.739 | 1
| 1 | 2 | 10.233 | 11.374 | 10.892 | 9.986 | 10.362 | 11.0297 | 11.06 | 10.7 | 8.4721 | 9.751 | 10.94 | 13.896 | 10.304 | 11.53 | 10.398 | 2
| 1 | 3 | 10.656 | 10.986 | 10.456 | 9.962 | 10.492 | 10.9009 | 10.82 | 10.82 | 8.4168 | 9.539 | 10.94 | 13.4846 | 10.551 | 11.76 | 10.833 | 16
| 1 | 4 | 10.127 | 10.139 | 10.386 | 10.586 | 10.574 | 10.9099 | 10.94 | 10.82 | 8.4646 | 9.539 | 10.94 | 13.6697 | 11.315 | 11.76 | 9.48 | 10.6743
| 1 | 5 | 10.409 | 11.28 | 10.48 | 10.304 | 10.656 | 11.0201 | 10.82 | 10.47 | 8.2418 | 9.739 | 10.82 | 13.4173 | 11.339 | 12.47 | 10.621 | 56
| 1 | 6 | 10.057 | 10.468 | 10.468 | 10.456 | 10.48 | 11.3113 | 10.82 | 10.7 | 8.5064 | 9.48 | 11.17 | 13.5786 | 10.997 | 12.23 | 11.009
| Mean |  | 10.2722 | 10.8055 | 10.4407 | 10.2505 | 10.4798 | 11.0427 | 10.92 | 10.7217 | 8.3859 | 9.6488 | 10.9 | 13.6468 | 10.8643 | 12.115 | 10.5133
| Median |  | 10.192 | 10.786 | 10.462 | 10.2565 | 10.486 | 11.0249 | 10.88 | 10.76 | 8.4407 | 9.639 | 10.94 | 13.6241 | 10.8385 | 11.995 | 10.68
| Std Dev. |  | 0.22344 | 0.48728 | 0.29609 | 0.25003 | 0.12736 | 0.14986 | 0.11798 | 0.13659 | 0.12617 | 0.14811 | 0.18984 | 0.19062 | 0.42241 | 0.53182 | 0.54616
| Rel.Std.Dev. |  | 2.175% | 4.51% | 2.836% | 2.439% | 1.215% | 1.357% | 1.08% | 1.274% | 1.505% | 1.535% | 1.742% | 1.397% | 3.888% | 4.39% | 5.195%
| PDM3 |  | -3.767% | 1.229% | -2.189% | -3.97% | -1.822% | 3.452% | 2.302% | 0.4439% | -21.44% | -9.607% | 2.115% | 27.85% | 1.78% | 13.5% | -1.508%
| Z-Score (Absolute) |  | 1.08 | 0.2 | 0.67 | 1.13 | 0.58 | 0.76 | 0.47 | 0 | 5.58 | 2.57 | 0.42 | 6.98 | 0.34 | 3.32 | 0.01
| Analytical results for Te in OREAS 461 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 1 | < 1 | [13 empty] | 1
| 1 | 2 | < 1 | < 1 | [13 empty] | 16
| 1 | 3 | < 1 | < 1 | [13 empty] | 16
| 1 | 4 | < 1 | < 1 | [13 empty] | < 1
| 1 | 5 | < 1 | < 1 | [13 empty] | 22
| 1 | 6 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for ThO2 in OREAS 461 (Certified Value 239 ppm) | [16 empty] | Certified
ThO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 228.72 | 233.27 | 240.1 | 232.13 | 234.41 | 237.0464 | 252.27 | 239.76 | 178.9585 | 174.9 | 230.42 | 170.7405 | 177.51 | 251.5 | 232.1 | 1
| 1 | 2 | 228.72 | 250.34 | 240.1 | 226.44 | 244.65 | 243.0596 | 250.11 | 238.96 | 177.9416 | 176.03 | 232.7 | 166.107 | 169.55 | 250.3 | 232.1 | #N/A
| 1 | 3 | 234.41 | 244.65 | 243.51 | 230.99 | 237.82 | 239.7084 | 250.11 | 238.39 | 178.5566 | 173.19 | 231.56 | 164.9048 | 170.68 | 251.5 | 219.6 | 16
| 1 | 4 | 229.86 | 227.58 | 237.82 | 235.55 | 246.92 | 239.9244 | 249.66 | 241.12 | 176.6683 | 171.37 | 233.27 | 167.2306 | 178.65 | 250.3 | 229.9 | 238.7942
| 1 | 5 | 230.99 | 245.79 | 238.96 | 235.55 | 241.23 | 236.0327 | 252.27 | 235.32 | 175.172 | 172.85 | 232.13 | 165.3768 | 178.65 | 252.6 | 232.1 | 57
| 1 | 6 | 227.58 | 230.99 | 240.1 | 237.82 | 245.79 | 239.6234 | 245.9 | 235.09 | 177.1456 | 170.34 | 241.23 | 166.8963 | 177.51 | 254.9 | 232.1
| Mean |  | 230.0467 | 238.77 | 240.0983 | 233.08 | 241.8033 | 239.2325 | 250.0533 | 238.1067 | 177.4071 | 173.1133 | 233.5517 | 166.876 | 175.425 | 251.85 | 229.65
| Median |  | 229.29 | 238.96 | 240.1 | 233.84 | 242.94 | 239.6659 | 250.11 | 238.675 | 177.5436 | 173.02 | 232.415 | 166.5017 | 177.51 | 251.5 | 232.1
| Std Dev. |  | 2.4318 | 9.3136 | 1.904 | 4.0995 | 4.92 | 2.4707 | 2.334 | 2.4285 | 1.3874 | 2.1216 | 3.8873 | 2.0873 | 4.16 | 1.7271 | 5.0015
| Rel.Std.Dev. |  | 1.057% | 3.901% | 0.793% | 1.759% | 2.035% | 1.033% | 0.9334% | 1.02% | 0.782% | 1.226% | 1.664% | 1.251% | 2.371% | 0.6858% | 2.178%
| PDM3 |  | -3.663% | -0.01014% | 0.5461% | -2.393% | 1.26% | 0.1835% | 4.715% | -0.2879% | -25.71% | -27.51% | -2.195% | -30.12% | -26.54% | 5.467% | -3.829%
| Z-Score (Absolute) |  | 0.29 | 0.55 | 0.67 | 0 | 0.84 | 0.59 | 1.63 | 0.48 | 5.35 | 5.76 | 0.1 | 6.36 | 5.54 | 1.8 | 0.13
| Analytical results for TiO2 in OREAS 461 (Certified Value 3.07 wt.%) | [16 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 3.106 | 3.219 | 3.019 | 3.1153 | 2.89 | 3.136 | 2.9645 | 2.953 | 3.1911 | 3.169 | 3.053 |  |  |  |  | 1
| 1 | 2 | 3.1327 | 3.169 | 3.019 | 3.0032 | 2.98 | 3.103 | 2.953 | 2.869 | 3.106 | 3.169 | 3.053 |  |  |  |  | 3
| 1 | 3 | 3.0893 | 3.186 | 3.003 | 3.1027 | 2.89 | 3.136 | 2.9606 | 2.669 | 3.0059 | 3.119 | 2.786 |  |  |  |  | 16
| 1 | 4 | 3.0793 | 3.203 | 3.019 | 3.0666 | 2.95 | 3.119 | 2.9808 | 2.285 | 3.0476 | 3.136 | 3.069 |  |  |  |  | 3.0692
| 1 | 5 | 3.0042 | 3.169 | 3.036 | 3.055 | 2.86 | 3.086 | 2.9462 | 2.852 | 3.111 | 3.186 | 3.119 |  |  |  |  | 58
| 1 | 6 | 3.121 | 3.119 | 3.003 | 3.0519 | 2.85 | 3.136 | 2.9923 | 2.919 | 3.2094 | 3.236 | 3.069
| Mean |  | 3.0888 | 3.1775 | 3.0165 | 3.0658 | 2.9033 | 3.1193 | 2.9662 | 2.7578 | 3.1118 | 3.1692 | 3.0248
| Median |  | 3.0976 | 3.1775 | 3.019 | 3.0608 | 2.89 | 3.1275 | 2.9626 | 2.8605 | 3.1085 | 3.169 | 3.061
| Std Dev. |  | 0.045842 | 0.034674 | 0.012357 | 0.040102 | 0.051251 | 0.02103 | 0.017338 | 0.2517 | 0.078975 | 0.040907 | 0.1195
| Rel.Std.Dev. |  | 1.484% | 1.091% | 0.4097% | 1.308% | 1.765% | 0.6742% | 0.5845% | 9.127% | 2.538% | 1.291% | 3.951%
| PDM3 |  | 0.6385% | 3.53% | -1.716% | -0.1104% | -5.403% | 1.635% | -3.353% | -10.14% | 1.391% | 3.259% | -1.444%
| Z-Score (Absolute) |  | 0.19 | 1.26 | 0.67 | 0.08 | 2.03 | 0.56 | 1.28 | 2.64 | 0.47 | 1.16 | 0
| Analytical results for Tl in OREAS 461 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 14 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 17
| 1 | 3 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | < 0.5
| 1 | 5 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 23
| 1 | 6 | < 0.5 | < 0.5 | 0.1 | < 0.5 | < 2 | < 0.5
| Mean |  |  |  | 0.1
| Median |  |  |  | 0.1
| Std Dev. |  |  |  | 1.6997e-17
| Rel.Std.Dev. |  |  |  | 1.7e-14%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Tm2O3 in OREAS 461 (Certified Value 1.02 ppm) | [16 empty] | Certified
Tm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 0.948 | 0.994 | 0.994 | 1.096 | 1.051 | 0.95011 | 1.14 | 1.14 | 0.62325 | 0.822 | 1.03 | 0.81389 | 1.028 | 1.6 | 0.937 | 1
| 1 | 2 | 0.959 | 1.051 | 0.982 | 1.005 | 1.016 | 0.91955 | 1.03 | 1.14 | 0.71153 | 0.765 | 1.03 | 0.84746 | 0.959 | 1.14 | 1.051 | 4
| 1 | 3 | 0.982 | 1.154 | 0.914 | 1.062 | 1.039 | 0.87755 | 1.03 | 1.14 | 0.66626 | 0.777 | 1.03 | 0.80106 | 1.039 | 1.37 | 1.051 | 16
| 1 | 4 | 0.959 | 1.028 | 1.016 | 1.051 | 1.005 | 1.0191 | 1.03 | 1.14 | 0.6406 | 0.777 | 1.03 | 0.79412 | 1.279 | 1.14 | 0.868 | 1.0158
| 1 | 5 | 0.959 | 1.131 | 1.039 | 1.108 | 1.005 | 0.96162 | 1.14 | 1.03 | 0.71229 | 0.799 | 1.03 | 0.76918 | 1.051 | 1.14 | 0.811 | 59
| 1 | 6 | 0.959 | 0.994 | 0.937 | 1.154 | 1.028 | 1.0205 | 1.03 | 1.14 | 0.63683 | 0.765 | 1.14 | 0.78947 | 1.051 | 1.14 | 0.914
| Mean |  | 0.961 | 1.0587 | 0.98033 | 1.0793 | 1.024 | 0.95808 | 1.0667 | 1.1217 | 0.66513 | 0.78417 | 1.0483 | 0.80253 | 1.0678 | 1.255 | 0.93867
| Median |  | 0.959 | 1.0395 | 0.988 | 1.079 | 1.022 | 0.95586 | 1.03 | 1.14 | 0.65343 | 0.777 | 1.03 | 0.79759 | 1.045 | 1.14 | 0.9255
| Std Dev. |  | 0.011189 | 0.068823 | 0.047289 | 0.051582 | 0.01874 | 0.055986 | 0.056804 | 0.044907 | 0.038818 | 0.022311 | 0.044907 | 0.026471 | 0.10901 | 0.19243 | 0.097099
| Rel.Std.Dev. |  | 1.164% | 6.501% | 4.824% | 4.779% | 1.83% | 5.844% | 5.325% | 4.004% | 5.836% | 2.845% | 4.284% | 3.298% | 10.21% | 15.33% | 10.34%
| PDM3 |  | -5.399% | 4.215% | -3.496% | 6.25% | 0.8028% | -5.686% | 5.003% | 10.42% | -34.52% | -22.81% | 3.198% | -21% | 5.118% | 23.54% | -7.597%
| Z-Score (Absolute) |  | 0.67 | 0.37 | 0.47 | 0.59 | 0 | 0.71 | 0.45 | 1.06 | 3.85 | 2.57 | 0.25 | 2.38 | 0.02 | 1.27 | 0.92
| Analytical results for U3O8 in OREAS 461 (Certified Value 5.65 ppm) | [16 empty] | Certified
U3O8 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 5.283 | 5.507 | 5.54 | 5.436 | 5.495 | 5.9366 | 5.66 | 5.78 | 4.4392 | 4.245 | 5.31 | 4.6899 | 3.62 | 5.9 | 5.931 | 1
| 1 | 2 | 5.259 | 5.943 | 5.42 | 5.424 | 5.519 | 5.6521 | 5.66 | 5.78 | 4.5209 | 4.304 | 5.31 | 4.7668 | 3.455 | 5.31 | 5.613 | 5
| 1 | 3 | 5.542 | 5.778 | 5.54 | 5.566 | 5.696 | 5.95 | 5.66 | 5.78 | 4.4971 | 4.174 | 5.31 | 4.6137 | 3.549 | 5.31 | 5.967 | 16
| 1 | 4 | 5.365 | 5.424 | 5.42 | 5.53 | 5.861 | 5.7731 | 5.66 | 5.66 | 4.2753 | 4.092 | 6.49 | 4.7793 | 3.797 | 5.9 | 6.038 | 5.6463
| 1 | 5 | 5.424 | 5.79 | 5.78 | 5.542 | 5.719 | 5.3888 | 5.66 | 5.78 | 4.571 | 4.056 | 5.31 | 4.7524 | 3.679 | 5.9 | 6.203 | 60
| 1 | 6 | 5.33 | 5.424 | 5.78 | 5.696 | 5.825 | 5.7676 | 5.54 | 5.78 | 4.5307 | 4.033 | 5.9 | 4.8462 | 3.585 | 5.9 | 6.014
| Mean |  | 5.3672 | 5.6443 | 5.58 | 5.5323 | 5.6858 | 5.7447 | 5.64 | 5.76 | 4.4724 | 4.1507 | 5.605 | 4.7414 | 3.6142 | 5.7033 | 5.961
| Median |  | 5.3475 | 5.6425 | 5.54 | 5.536 | 5.7075 | 5.7703 | 5.66 | 5.78 | 4.509 | 4.133 | 5.31 | 4.7596 | 3.6025 | 5.9 | 5.9905
| Std Dev. |  | 0.10389 | 0.22102 | 0.16395 | 0.099031 | 0.15199 | 0.20763 | 0.04899 | 0.04899 | 0.10585 | 0.10881 | 0.49363 | 0.080188 | 0.1167 | 0.30467 | 0.1946
| Rel.Std.Dev. |  | 1.936% | 3.916% | 2.938% | 1.79% | 2.673% | 3.614% | 0.8686% | 0.8505% | 2.367% | 2.622% | 8.807% | 1.691% | 3.229% | 5.342% | 3.265%
| PDM3 |  | -4.944% | -0.03482% | -1.174% | -2.018% | 0.7002% | 1.743% | -0.1116% | 2.014% | -20.79% | -26.49% | -0.7314% | -16.03% | -35.99% | 1.01% | 5.574%
| Z-Score (Absolute) |  | 0.89 | 0.26 | 0 | 0.2 | 0.43 | 0.67 | 0.24 | 0.73 | 4.42 | 5.91 | 0.65 | 3.47 | 8.13 | 0.49 | 1.86
| Analytical results for V2O5 in OREAS 461 (Certified Value 687 ppm) | [16 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 681.9 | 681.9 | 696.2 | 699.8 | 710.5 | 733.7 | 669.4 | 660.5 | 710 | 647 | 664.1 | 714 | 664.1 |  |  | 1
| 1 | 2 | 680.2 | 749.8 | 665.9 | 676.6 | 701.6 | 723 | 676.6 | 683.7 | 710 | 643 | 667.7 | 714 | 657 |  |  | 6
| 1 | 3 | 685.5 | 703.4 | 676.6 | 699.8 | 737.3 | 708.7 | 676.6 | 633.7 | 710 | 649 | 678.4 | 714 | 671.2 |  |  | 16
| 1 | 4 | 667.7 | 699.8 | 689.1 | 694.4 | 717.7 | 735.5 | 664.1 | 632 | 710 | 661 | 660.5 | 714 | 664.1 |  |  | 686.6225
| 1 | 5 | 674.8 | 733.7 | 710.5 | 701.6 | 703.4 | 721.2 | 671.2 | 614.1 | 710 | 651 | 633.7 | 714 | 667.7 |  |  | 61
| 1 | 6 | 669.4 | 708.7 | 689.1 | 705.2 | 719.4 | 714.1 | 648 | 649.8 | 710 | 652 | 665.9 | 625 | 674.8
| Mean |  | 676.5833 | 712.8833 | 687.9 | 696.2333 | 714.9833 | 722.7 | 667.65 | 645.6333 | 710 | 650.5 | 661.7167 | 699.1667 | 666.4833
| Median |  | 677.5 | 706.05 | 689.1 | 699.8 | 714.1 | 722.1 | 670.3 | 641.75 | 710 | 650 | 665 | 714 | 665.9
| Std Dev. |  | 7.1351 | 24.6363 | 15.4546 | 10.232 | 13.102 | 10.5559 | 10.72 | 24.536 | 0 | 6.0581 | 14.9915 | 36.3341 | 6.232
| Rel.Std.Dev. |  | 1.055% | 3.456% | 2.247% | 1.47% | 1.832% | 1.461% | 1.606% | 3.8% | 0% | 0.9313% | 2.266% | 5.197% | 0.9351%
| PDM3 |  | -1.462% | 3.825% | 0.1861% | 1.4% | 4.13% | 5.254% | -2.763% | -5.97% | 3.405% | -5.261% | -3.627% | 1.827% | -2.933%
| Z-Score (Absolute) |  | 0.36 | 0.79 | 0 | 0.39 | 0.85 | 1.1 | 0.64 | 1.33 | 0.82 | 1.18 | 0.65 | 0.36 | 0.67
| Analytical results for WO3 in OREAS 461 (Certified Value 4.89 ppm) | [16 empty] | Certified
WO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 17 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 5 | 5 | 5 | 3.8 | 5 | 5 | 3.8 | < 13 | 3.8 | < 6.3 | 5 |  |  |  |  | 1
| 1 | 2 | 5 | 5 | 5 | 3.8 | 5 | 5 | 3.8 | < 13 | 3.8 | < 6.3 | 2.5 |  |  |  |  | 7
| 1 | 3 | 6.3 | 5 | 5 | 5 | 5 | 5 | 3.8 | 25 | 3.8 | < 6.3 | 13.9 |  |  |  |  | 16
| 1 | 4 | 5 | 5 | 5 | 5 | 5 | 5 | 3.8 | < 13 | 2.5 | < 6.3 | 12.6 |  |  |  |  | 4.8899
| 1 | 5 | 8.8 | 6.3 | 6.3 | 3.8 | 5 | 5 | 3.8 | < 13 | 3.8 | 6.3 | 10.1 |  |  |  |  | 62
| 1 | 6 | 6.3 | 5 | 5 | 5 | 5 | 3.8 | 3.8 | 13 | 3.8 | < 6.3 | 11.3
| Mean |  | 6.0667 | 5.2167 | 5.2167 | 4.4 | 5 | 4.8 | 3.8 | 19 | 3.5833 | 6.3 | 9.2333
| Median |  | 5.65 | 5 | 5 | 4.4 | 5 | 5 | 3.8 | 19 | 3.8 | 6.3 | 10.7
| Std Dev. |  | 1.4828 | 0.53072 | 0.53072 | 0.65727 | 0 | 0.4899 | 4.8648e-16 | 8.4853 | 0.53072 |  | 4.5032
| Rel.Std.Dev. |  | 24.44% | 10.17% | 10.17% | 14.94% | 0% | 10.21% | 1.28e-14% | 44.66% | 14.81% |  | 48.77%
| PDM3 |  | 24.07% | 6.683% | 6.683% | -10.02% | 2.252% | -1.838% | -22.29% | 288.6% | -26.72% | 28.84% | 88.83%
| Z-Score (Absolute) |  | 0.24 | 0 | 0 | 0.67 | 0.17 | 0.34 | 1.18 | 10.96 | 1.35 | 0.84 | 3.2
| Analytical results for Y2O3 in OREAS 461 (Certified Value 116 ppm) | [16 empty] | Certified
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 115.81 | 119.88 | 119.4 | 117.47 | 109.08 | 112.038 | 121.91 | 117.21 | 93.592 | 108.58 | 100.5069 | 194.29 | 115.6 | 111.75 |  | 1
| 1 | 2 | 113.15 | 129.53 | 121.9 | 114.16 | 109.21 | 113.6094 | 124.83 | 116.83 | 91.814 | 109.21 | 104.8175 | 187.95 | 115.6 | 110.86 |  | 8
| 1 | 3 | 118.1 | 123.56 | 119.4 | 117.59 | 111.88 | 114.2833 | 120.13 | 117.21 | 89.337 | 111.75 | 102.0242 | 185.41 | 116.8 | 93.21 |  | 16
| 1 | 4 | 113.4 | 122.8 | 124.5 | 117.72 | 110.99 | 114.6675 | 120.39 | 117.97 | 90.925 | 109.85 | 105.3273 | 157.47 | 114.3 | 112.13 |  | 115.5417
| 1 | 5 | 113.4 | 128.26 | 121.9 | 117.97 | 110.35 | 114.3567 | 121.53 | 115.56 | 90.633 | 110.48 | 109.2084 | 194.29 | 119.4 | 113.66 |  | 63
| 1 | 6 | 113.66 | 121.78 | 121.9 | 118.48 | 110.74 | 114.999 | 117.72 | 117.21 | 89.833 | 115.56 | 115.0991 | 186.68 | 119.4 | 114.29
| Mean |  | 114.5867 | 124.3017 | 121.5 | 117.2317 | 110.375 | 113.9923 | 121.085 | 116.9983 | 91.0223 | 110.905 | 106.1639 | 184.3483 | 116.85 | 109.3167
| Median |  | 113.53 | 123.18 | 121.9 | 117.655 | 110.545 | 114.32 | 120.96 | 117.21 | 90.779 | 110.165 | 105.0724 | 187.315 | 116.2 | 111.94
| Std Dev. |  | 1.9788 | 3.7873 | 1.9131 | 1.5468 | 1.0781 | 1.0633 | 2.3504 | 0.79695 | 1.5262 | 2.5281 | 5.3058 | 13.7088 | 2.1277 | 7.9905
| Rel.Std.Dev. |  | 1.727% | 3.047% | 1.575% | 1.319% | 0.9767% | 0.9327% | 1.941% | 0.6812% | 1.677% | 2.279% | 4.998% | 7.436% | 1.821% | 7.31%
| PDM3 |  | -0.8266% | 7.582% | 5.157% | 1.463% | -4.472% | -1.341% | 4.798% | 1.261% | -21.22% | -4.013% | -8.116% | 59.55% | 1.132% | -5.388%
| Z-Score (Absolute) |  | 0.15 | 1.14 | 0.77 | 0.2 | 0.71 | 0.23 | 0.72 | 0.17 | 3.28 | 0.64 | 1.27 | 9.85 | 0.15 | 0.42
| Analytical results for Yb2O3 in OREAS 461 (Certified Value 5 ppm) | [16 empty] | Certified
Yb2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.703 | 4.874 | 5.272 | 4.68 | 4.931 | 5.0147 | 5.24 | 5.01 | 4.0939 | 3.99 | 4.67 | 4.8668 | 6.15 | 4.55 | 5.47 | 1
| 1 | 2 | 4.418 | 5.523 | 5.329 | 4.805 | 4.646 | 4.8817 | 5.35 | 5.12 | 3.8286 | 3.99 | 4.78 | 4.8461 | 5.92 | 5.12 | 5.35 | 9
| 1 | 3 | 5.101 | 5.329 | 5.215 | 4.635 | 4.805 | 4.6217 | 5.12 | 4.9 | 4.0425 | 3.76 | 5.12 | 4.7279 | 5.69 | 5.69 | 6.38 | 16
| 1 | 4 | 4.794 | 4.908 | 4.908 | 4.896 | 4.919 | 4.8367 | 5.12 | 5.24 | 4.1623 | 3.76 | 5.01 | 4.7017 | 5.01 | 5.12 | 5.69 | 5.0035
| 1 | 5 | 4.714 | 4.999 | 5.466 | 4.828 | 5.033 | 4.9337 | 5.24 | 5.12 | 4.2906 | 3.87 | 5.12 | 4.4511 | 6.26 | 5.12 | 5.12 | 64
| 1 | 6 | 4.703 | 5.192 | 4.737 | 4.942 | 4.76 | 4.8863 | 5.12 | 5.24 | 4.004 | 3.76 | 5.01 | 4.6117 | 5.92 | 5.12 | 5.58
| Mean |  | 4.7388 | 5.1375 | 5.1545 | 4.7977 | 4.849 | 4.8625 | 5.1983 | 5.105 | 4.0703 | 3.855 | 4.9517 | 4.7009 | 5.825 | 5.12 | 5.5983
| Median |  | 4.7085 | 5.0955 | 5.2435 | 4.8165 | 4.862 | 4.884 | 5.18 | 5.12 | 4.0682 | 3.815 | 5.01 | 4.7148 | 5.92 | 5.12 | 5.525
| Std Dev. |  | 0.21924 | 0.25672 | 0.27569 | 0.11982 | 0.13896 | 0.13252 | 0.094745 | 0.13263 | 0.15567 | 0.11292 | 0.18563 | 0.15471 | 0.44617 | 0.3605 | 0.43051
| Rel.Std.Dev. |  | 4.626% | 4.997% | 5.348% | 2.497% | 2.866% | 2.725% | 1.823% | 2.598% | 3.825% | 2.929% | 3.749% | 3.291% | 7.66% | 7.041% | 7.69%
| PDM3 |  | -5.29% | 2.678% | 3.017% | -4.114% | -3.088% | -2.819% | 3.894% | 2.028% | -18.65% | -22.95% | -1.036% | -6.049% | 16.42% | 2.328% | 11.89%
| Z-Score (Absolute) |  | 0.96 | 0.62 | 0.67 | 0.52 | 0.35 | 0.14 | 0.83 | 0.51 | 2.96 | 3.69 | 0 | 0.85 | 3.47 | 0.57 | 1.64
| Analytical results for ZnO in OREAS 461 (Certified Value 201 ppm) | [16 empty] | Certified
ZnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 14 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 178 | 250 | 301.7375 | 237 | 187 | 171.8 | [9 empty] | 1
| 1 | 2 | 189.2 | 250 | 313.2752 | 212 | 187 | 164.3 | [9 empty] | #N/A
| 1 | 3 | 181.7 | 250 | 319.1172 | 212 | 187 | 171.8 | [9 empty] | 16
| 1 | 4 | 190.5 | 250 | 307.5917 | 212 | 187 | 173 | [9 empty] | 201.3257
| 1 | 5 | 194.2 | 250 | 292.2086 | 224 | 187 | 164.3 | [9 empty] | 65
| 1 | 6 | 196.7 | 250 | 322.7746 | 212 | 187 | 165.6
| Mean |  | 188.3833 | 250 | 309.4508 | 218.1667 | 187 | 168.4667
| Median |  | 189.85 | 250 | 310.4335 | 212 | 187 | 168.7
| Std Dev. |  | 7.2198 | 0 | 11.3652 | 10.4003 | 0 | 4.1404
| Rel.Std.Dev. |  | 3.833% | 0% | 3.673% | 4.767% | 0% | 2.458%
| PDM3 |  | -6.429% | 24.18% | 53.71% | 8.365% | -7.116% | -16.32%
| Z-Score (Absolute) |  | 0.37 | 1.36 | 3.09 | 0.37 | 0.41 | 0.93
| Analytical results for ZrO2 in OREAS 461 (Certified Value 815 ppm) | [16 empty] | Certified
ZrO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 805.1 | 852.4 | 829.4 | 830.7 | 790.2 | 836.1 | 810.5 | 792.9 | 757.8 | 784 | 888.83 |  |  |  |  | 1
| 1 | 2 | 815.9 | 922.6 | 844.2 | 821.3 | 809.1 | 838.8 | 782.1 | 844.2 | 778.1 | 792 | 887.48 |  |  |  |  | 10
| 1 | 3 | 819.9 | 880.7 | 847 | 824 | 817.2 | 828 | 786.2 | 788.9 | 768.6 | 762 | 860.46 |  |  |  |  | 16
| 1 | 4 | 792.9 | 860.5 | 857.8 | 847 | 842.9 | 822.6 | 794.3 | 837.5 | 763.2 | 785 | 725.38 |  |  |  |  | 814.5465
| 1 | 5 | 774 | 888.8 | 879.4 | 821.3 | 833.4 | 822.6 | 794.3 | 790.2 | 745.6 | 767 | 918.54 |  |  |  |  | 66
| 1 | 6 | 784.8 | 841.5 | 856.4 | 849.7 | 802.4 | 822.6 | 813.2 | 795.6 | 763.2 | 763 | 879.37
| Mean |  | 798.7667 | 874.4167 | 852.3667 | 832.3333 | 815.8667 | 828.45 | 796.7667 | 808.2167 | 762.75 | 775.5 | 860.01
| Median |  | 799 | 870.6 | 851.7 | 827.35 | 813.15 | 825.3 | 794.3 | 794.25 | 763.2 | 775.5 | 883.425
| Std Dev. |  | 18.0175 | 29.4276 | 16.7305 | 12.9019 | 19.6261 | 7.3282 | 12.6301 | 25.4709 | 10.8524 | 13.0038 | 68.5729
| Rel.Std.Dev. |  | 2.256% | 3.365% | 1.963% | 1.55% | 2.406% | 0.8846% | 1.585% | 3.151% | 1.423% | 1.677% | 7.973%
| PDM3 |  | -1.937% | 7.35% | 4.643% | 2.184% | 0.1621% | 1.707% | -2.183% | -0.7771% | -6.359% | -4.794% | 5.581%
| Z-Score (Absolute) |  | 0.6 | 2.06 | 1.29 | 0.58 | 0 | 0.44 | 0.67 | 0.27 | 1.87 | 1.42 | 2.5
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 461 (Certified Value < 2 ppm) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 0.59 | 1.01 | < 0.01 | 0.18 | 0.26 | 0.47462 | 0.6 | 0.9 | < 0.01 | 1.46 | < 1 | 1 |  |  | 1
| 1 | 2 | 0.58 | 1.14 | 0.86 | 0.22 | 0.21 | 0.44091 | 0.8 | 0.9 | < 0.01 | 1.02 | < 1 | 0.5 |  |  | 15
| 1 | 3 | 0.33 | 1.07 | < 0.01 | 0.13 | 0.33 | 0.41447 | 0.7 | 1 | < 0.01 | 1.56 | < 1 | 0.5 |  |  | 16
| 1 | 4 | 0.71 | 1.2 | 1.86 | < 0.01 | 0.33 | 0.4622 | 0.6 | 0.7 | < 0.01 | 1.43 | < 1 | 1 |  |  | < 2
| 1 | 5 | 0.09 | 0.91 | < 0.01 | 0.48 | 0.38 | 0.50033 | 0.9 | 1.5 | < 0.01 | 1.25 | < 1 | 1 |  |  | 68
| 1 | 6 | 0.96 | 1 | 1.4 | 0.09 | 0.24 | 0.50642 | 0.7 | 1.3 | < 0.01 | 1.49 | < 1 | 1
| Mean |  | 0.54333 | 1.055 | 1.3733 | 0.22 | 0.29167 | 0.46649 | 0.71667 | 1.05 |  | 1.3683 |  | 0.83333
| Median |  | 0.585 | 1.04 | 1.4 | 0.18 | 0.295 | 0.46841 | 0.7 | 0.95 |  | 1.445 |  | 1
| Std Dev. |  | 0.30197 | 0.10445 | 0.50053 | 0.15346 | 0.064936 | 0.035169 | 0.1169 | 0.29496 |  | 0.19954 |  | 0.2582
| Rel.Std.Dev. |  | 55.58% | 9.901% | 36.45% | 69.75% | 22.26% | 7.539% | 16.31% | 28.09% |  | 14.58% |  | 30.98%
PDM3
| Z-Score (Absolute) |  | 0.53 | 0.64 | 1.37 | 1.42 | 1.11 | 0.71 | 0.13 | 0.63 | NA | 1.52 | NA | 0.13
| Analytical results for Al in OREAS 461 (Certified Value 5.96 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 5.11 | 5.98 | 5.73 | 6.22 | 6.15 | 5.7075 | 5.95 | 5.86 | 6.02 | 5.8334 | 5.9812 | 5.94 | 6.13 | 6.19 | 1
| 1 | 2 | 5.69 | 6.05 | 5.73 | 6.15 | 6.09 | 5.5674 | 5.87 | 5.84 | 6.01 | 5.9219 | 6.0334 | 5.72 | 6.142 | 6.17 | #N/A
| 1 | 3 | 5.36 | 6.1 | 5.71 | 6.23 | 5.79 | 5.6788 | 5.85 | 5.94 | 6.01 | 5.8959 | 5.9544 | 5.85 | 6.119 | 6.27 | 16
| 1 | 4 | 5.52 | 6.07 | 5.76 | 6.18 | 5.64 | 5.937 | 5.93 | 5.88 | 6.04 | 6.0783 | 5.9523 | 5.99 | 6.124 | 6.09 | 5.9585
| 1 | 5 | 5.03 | 6.1 | 5.72 | 6.33 | 5.97 | 5.7769 | 5.91 | 5.91 | 5.97 | 5.8137 | 5.8569 | 5.83 | 6.177 | 6.16 | 69
| 1 | 6 | 4.94 | 6.13 | 5.72 | 6.29 | 5.64 | 5.6531 | 5.91 | 6.01 | 5.99 | 5.8489 | 5.854 | 5.86 | 6.115 | 6.38
| Mean |  | 5.275 | 6.0717 | 5.7283 | 6.2333 | 5.88 | 5.7201 | 5.9033 | 5.9067 | 6.0067 | 5.8987 | 5.9387 | 5.865 | 6.1345 | 6.21
| Median |  | 5.235 | 6.085 | 5.725 | 6.225 | 5.88 | 5.6932 | 5.91 | 5.895 | 6.01 | 5.8724 | 5.9533 | 5.855 | 6.127 | 6.18
| Std Dev. |  | 0.29629 | 0.052694 | 0.017224 | 0.067132 | 0.22289 | 0.12643 | 0.037238 | 0.06186 | 0.024221 | 0.096717 | 0.070804 | 0.093541 | 0.022863 | 0.10139
| Rel.Std.Dev. |  | 5.617% | 0.8679% | 0.3007% | 1.077% | 3.791% | 2.21% | 0.6308% | 1.047% | 0.4032% | 1.64% | 1.192% | 1.595% | 0.3727% | 1.633%
| PDM3 |  | -11.47% | 1.898% | -3.864% | 4.612% | -1.318% | -4.001% | -0.9266% | -0.8707% | 0.8076% | -1.005% | -0.3329% | -1.57% | 2.953% | 4.22%
| Z-Score (Absolute) |  | 3.17 | 0.84 | 0.89 | 1.65 | 0.13 | 0.93 | 0.01 | 0.01 | 0.51 | 0.21 | 0.17 | 0.2 | 1.15 | 1.53
| Analytical results for As in OREAS 461 (Certified Value 33.6 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 34 | 34.8 | 34.2 | 33 | 33.2 | 30.5833 | 34 | 34 | 33 | 48.2 | < 50 | 35 |  |  | 1
| 1 | 2 | 33.5 | 32.5 | 34.6 | 32.7 | 31.9 | 27.5692 | 36 | 32 | 33.2 | 44.1 | < 50 | 34 |  |  | 16
| 1 | 3 | 33.8 | 33.4 | 35.4 | 32.5 | 33 | 28.2683 | 35 | 35 | 32.5 | 49.7 | < 50 | 37 |  |  | 16
| 1 | 4 | 34.8 | 33.3 | 32.3 | 32.5 | 32 | 30.45 | 35 | 34 | 32.4 | 51.1 | < 50 | 34 |  |  | 33.5537
| 1 | 5 | 34 | 32.8 | 34.2 | 32.8 | 33.1 | 29.5769 | 33 | 33 | 32 | 50.4 | < 50 | 35 |  |  | 70
| 1 | 6 | 33.9 | 32.3 | 32.6 | 32.2 | 33.2 | 29.8258 | 33 | 36 | 32.7 | 50.5 | < 50 | 35
| Mean |  | 34 | 33.1833 | 33.8833 | 32.6167 | 32.7333 | 29.3789 | 34.3333 | 34 | 32.6333 | 49 |  | 35
| Median |  | 33.95 | 33.05 | 34.2 | 32.6 | 33.05 | 29.7014 | 34.5 | 34 | 32.6 | 50.05 |  | 35
| Std Dev. |  | 0.43359 | 0.90203 | 1.1974 | 0.27869 | 0.6121 | 1.212 | 1.2111 | 1.4142 | 0.43205 | 2.5985 |  | 1.0954
| Rel.Std.Dev. |  | 1.275% | 2.718% | 3.534% | 0.8544% | 1.87% | 4.126% | 3.527% | 4.159% | 1.324% | 5.303% |  | 3.13%
| PDM3 |  | 1.33% | -1.104% | 0.9824% | -2.793% | -2.445% | -12.44% | 2.324% | 1.33% | -2.743% | 46.03% |  | 4.31%
| Z-Score (Absolute) |  | 0.11 | 0.66 | 0 | 1.19 | 1.08 | 4.24 | 0.42 | 0.11 | 1.18 | 15.15 | NA | 0.67
| Analytical results for Ba in OREAS 461 (Certified Value 957 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 930 | 950 | 980 | 1020 | 1000 | 930.1279 | 937 | 913 | 1020 | 945.2 | 1030.99 | 330 | 945 | 918 | 1
| 1 | 2 | 950 | 950 | 970 | 1000 | 1000 | 926.4228 | 945 | 928 | 1040 | 952.3 | 1026.01 | 334 | 955 | 890 | #N/A
| 1 | 3 | 880 | 950 | 970 | 1030 | 950 | 940.008 | 970 | 916 | 1050 | 952.1 | 1005.06 | 348 | 950 | 896 | 16
| 1 | 4 | 890 | 940 | 960 | 1010 | 940 | 926.149 | 945 | 911 | 1030 | 968.7 | 1018.382 | 335 | 955 | 898 | 957.0984
| 1 | 5 | 870 | 940 | 960 | 1030 | 1000 | 949.4869 | 931 | 921 | 1020 | 922.5 | 1009.25 | 350 | 945 | 899 | 71
| 1 | 6 | 880 | 950 | 970 | 1020 | 920 | 932.7722 | 932 | 932 | 1040 | 931.1 | 1004.92 | 352 | 940 | 950
| Mean |  | 900 | 946.6667 | 968.3333 | 1018.3333 | 968.3333 | 934.1611 | 943.3333 | 920.1667 | 1033.3333 | 945.3167 | 1015.7687 | 341.5 | 948.3333 | 908.5
| Median |  | 885 | 950 | 970 | 1020 | 975 | 931.45 | 941 | 918.5 | 1035 | 948.65 | 1013.816 | 341.5 | 947.5 | 898.5
| Std Dev. |  | 32.249 | 5.164 | 7.5277 | 11.6905 | 36.0093 | 9.0701 | 14.4037 | 8.4242 | 12.1106 | 16.5184 | 11.1199 | 9.5446 | 6.0553 | 22.4121
| Rel.Std.Dev. |  | 3.583% | 0.5455% | 0.7774% | 1.148% | 3.719% | 0.9709% | 1.527% | 0.9155% | 1.172% | 1.747% | 1.095% | 2.795% | 0.6385% | 2.467%
| PDM3 |  | -5.966% | -1.09% | 1.174% | 6.398% | 1.174% | -2.397% | -1.438% | -3.859% | 7.965% | -1.231% | 6.13% | -64.32% | -0.9158% | -5.078%
| Z-Score (Absolute) |  | 1.29 | 0.02 | 0.63 | 2.03 | 0.63 | 0.33 | 0.07 | 0.72 | 2.45 | 0.02 | 1.95 | 16.93 | 0.07 | 1.28
| Analytical results for Be in OREAS 461 (Certified Value 2.09 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 1.6 | 2.36 | 1.73 | 2.4 | 1.99 | 1.8332 | 2.3 | 2.4 | 2.04 | 2.16 | 1.9561 | 3 | 2.3 |  | 1
| 1 | 2 | 1.65 | 2.36 | 1.9 | 2.24 | 1.92 | 1.9227 | 2.3 | 1.5 | 2.01 | 2.26 | 2.2213 | 3 | 2.2 |  | #N/A
| 1 | 3 | 1.66 | 2.32 | 1.79 | 2.32 | 2.03 | 2.0916 | 2.4 | 1.9 | 2.04 | 2.33 | 2.0178 | 2.5 | 2.3 |  | 16
| 1 | 4 | 1.66 | 2.31 | 2.01 | 2.44 | 1.87 | 2.1006 | 2.3 | 1.9 | 2.02 | 2.32 | 2.0433 | 2.5 | 2.1 |  | 2.0913
| 1 | 5 | 1.77 | 2.14 | 1.83 | 2.36 | 1.9 | 2.1986 | 2.3 | 2.7 | 2 | 2.24 | 1.9839 | 3 | 2.1 |  | 72
| 1 | 6 | 1.72 | 2.14 | 1.83 | 2.5 | 1.87 | 2.2449 | 2.2 | 2.2 | 2.02 | 2.34 | 2.0687 | 2.5 | 2.3
| Mean |  | 1.6767 | 2.2717 | 1.8483 | 2.3767 | 1.93 | 2.0652 | 2.3 | 2.1 | 2.0217 | 2.275 | 2.0485 | 2.75 | 2.2167
| Median |  | 1.66 | 2.315 | 1.83 | 2.38 | 1.91 | 2.0961 | 2.3 | 2.05 | 2.02 | 2.29 | 2.0305 | 2.75 | 2.25
| Std Dev. |  | 0.059554 | 0.104 | 0.096833 | 0.091579 | 0.06603 | 0.15885 | 0.063246 | 0.42426 | 0.016021 | 0.06921 | 0.093758 | 0.27386 | 0.098319
| Rel.Std.Dev. |  | 3.552% | 4.578% | 5.239% | 3.853% | 3.421% | 7.692% | 2.75% | 20.2% | 0.7925% | 3.042% | 4.577% | 9.959% | 4.435%
| PDM3 |  | -19.83% | 8.623% | -11.62% | 13.64% | -7.714% | -1.247% | 9.978% | 0.4149% | -3.331% | 8.783% | -2.047% | 31.5% | 5.994%
| Z-Score (Absolute) |  | 1.66 | 0.67 | 0.99 | 1.09 | 0.67 | 0.14 | 0.79 | 0 | 0.31 | 0.69 | 0.34 | 2.55 | 0.46
| Analytical results for Bi in OREAS 461 (Certified Value 2.43 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.55 | 2.48 | 2.54 | 2.4 | 2.18 | 2.2786 | 2.5 | 2.6 | < 0.01 | 2.48 | 2.3108 | 2.4 | 2.3 |  | 1
| 1 | 2 | 2.56 | 2.48 | 2.58 | 2.39 | 2.14 | 2.1842 | 2.5 | 2.7 | < 0.01 | 2.49 | 2.0975 | 3.8 | 2.3 |  | #N/A
| 1 | 3 | 2.57 | 2.48 | 2.63 | 2.47 | 2.22 | 2.3718 | 2.6 | 2.6 | < 0.01 | 2.55 | 2.2453 | 2.2 | 2.4 |  | 16
| 1 | 4 | 2.58 | 2.55 | 2.57 | 2.52 | 2.12 | 2.4188 | 2.6 | 2.5 | < 0.01 | 2.58 | 2.0735 | 2.4 | 2.3 |  | 2.4269
| 1 | 5 | 2.51 | 2.53 | 2.62 | 2.48 | 2.17 | 2.4444 | 2.6 | 2.6 | < 0.01 | 2.46 | 2.1496 | 2.4 | 2.4 |  | 73
| 1 | 6 | 2.57 | 2.45 | 2.49 | 2.51 | 2.19 | 2.2143 | 2.4 | 2.7 | < 0.01 | 2.42 | 2.1767 | 2.4 | 2.5
| Mean |  | 2.5567 | 2.495 | 2.5717 | 2.4617 | 2.17 | 2.3187 | 2.5333 | 2.6167 |  | 2.4967 | 2.1756 | 2.6 | 2.3667
| Median |  | 2.565 | 2.48 | 2.575 | 2.475 | 2.175 | 2.3252 | 2.55 | 2.6 |  | 2.485 | 2.1632 | 2.4 | 2.35
| Std Dev. |  | 0.025033 | 0.037283 | 0.051929 | 0.054924 | 0.035777 | 0.10886 | 0.08165 | 0.075277 |  | 0.058878 | 0.089841 | 0.5933 | 0.08165
| Rel.Std.Dev. |  | 0.9791% | 1.494% | 2.019% | 2.231% | 1.649% | 4.695% | 3.223% | 2.877% |  | 2.358% | 4.129% | 22.82% | 3.45%
| PDM3 |  | 5.348% | 2.807% | 5.966% | 1.433% | -10.58% | -4.458% | 4.386% | 7.82% |  | 2.875% | -10.35% | 7.133% | -2.481%
| Z-Score (Absolute) |  | 0.52 | 0.11 | 0.61 | 0.11 | 2.03 | 1.05 | 0.36 | 0.91 | NA | 0.12 | 1.99 | 0.78 | 0.73
| Analytical results for Ca in OREAS 461 (Certified Value 1.21 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.23 | 1.25 | 1.06 | 1.14 | 1.07 | 1.155 | 1.22 | 1.29 | 1.16 | 1.3214 | 1.2546 | 1.18 | 1.2735 | 1.37 | 1
| 1 | 2 | 1.25 | 1.24 | 1.06 | 1.13 | 1.11 | 1.158 | 1.21 | 1.29 | 1.17 | 1.3087 | 1.2626 | 1.17 | 1.2725 | 1.37 | #N/A
| 1 | 3 | 1.21 | 1.26 | 1.04 | 1.15 | 1.07 | 1.1894 | 1.2 | 1.31 | 1.17 | 1.3334 | 1.254 | 1.2 | 1.272 | 1.4 | 16
| 1 | 4 | 1.23 | 1.29 | 1.06 | 1.14 | 1.02 | 1.2153 | 1.22 | 1.28 | 1.17 | 1.3683 | 1.2406 | 1.19 | 1.2715 | 1.37 | 1.2147
| 1 | 5 | 1.2 | 1.26 | 1.06 | 1.18 | 1.1 | 1.1753 | 1.22 | 1.29 | 1.16 | 1.268 | 1.2505 | 1.15 | 1.2675 | 1.38 | 74
| 1 | 6 | 1.21 | 1.26 | 1.06 | 1.16 | 1.02 | 1.1478 | 1.22 | 1.32 | 1.16 | 1.3018 | 1.2446 | 1.2 | 1.2685 | 1.4
| Mean |  | 1.2217 | 1.26 | 1.0567 | 1.15 | 1.065 | 1.1735 | 1.215 | 1.2967 | 1.165 | 1.3169 | 1.2511 | 1.1817 | 1.2709 | 1.3817
| Median |  | 1.22 | 1.26 | 1.06 | 1.145 | 1.07 | 1.1667 | 1.22 | 1.29 | 1.165 | 1.3151 | 1.2522 | 1.185 | 1.2717 | 1.375
| Std Dev. |  | 0.018348 | 0.016733 | 0.008165 | 0.017889 | 0.038341 | 0.025496 | 0.0083666 | 0.015055 | 0.0054772 | 0.033529 | 0.0078239 | 0.019408 | 0.0023752 | 0.01472
| Rel.Std.Dev. |  | 1.502% | 1.328% | 0.7727% | 1.556% | 3.6% | 2.173% | 0.6886% | 1.161% | 0.4701% | 2.546% | 0.6253% | 1.642% | 0.1869% | 1.065%
| PDM3 |  | 0.5737% | 3.729% | -13.01% | -5.326% | -12.32% | -3.395% | 0.02482% | 6.748% | -4.091% | 8.416% | 3% | -2.719% | 4.628% | 13.75%
| Z-Score (Absolute) |  | 0.04 | 0.53 | 2.06 | 0.87 | 1.95 | 0.57 | 0.04 | 1 | 0.68 | 1.26 | 0.42 | 0.47 | 0.67 | 2.08
| Analytical results for Cd in OREAS 461 (Certified Value < 0.1 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.04 | 0.05 | 0.07 | 0.04 | 0.07 | 0.07 | < 0.2 | 0.44 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 1
| 1 | 2 | 0.06 | < 0.02 | 0.08 | 0.07 | 0.05 | 0.076464 | < 0.2 | 0.41 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 18
| 1 | 3 | 0.08 | 0.04 | 0.06 | < 0.02 | 0.05 | 0.07 | < 0.2 | 0.4 | 0.03 | < 0.5 | < 0.5 |  |  |  | 16
| 1 | 4 | 0.04 | 0.03 | 0.09 | 0.05 | 0.05 | 0.07 | < 0.2 | 0.41 | < 0.02 | < 0.5 | < 0.5 |  |  |  | < 0.1
| 1 | 5 | 0.03 | 0.04 | 0.04 | 0.03 | 0.06 | 0.07 | 0.2 | 0.4 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 75
| 1 | 6 | 0.05 | 0.06 | 0.03 | 0.04 | 0.06 | 0.071069 | < 0.2 | 0.39 | < 0.02 | < 0.5 | < 0.5
| Mean |  | 0.05 | 0.044 | 0.061667 | 0.046 | 0.056667 | 0.071256 | 0.2 | 0.40833 | 0.03
| Median |  | 0.045 | 0.04 | 0.065 | 0.04 | 0.055 | 0.07 | 0.2 | 0.405 | 0.03
| Std Dev. |  | 01: 0.017889; 02: 0.011402; 03: 0.023166; 05: 0.015166; 06: 0.008165; 07: 0.0025873; 09: 0.017224
| Rel.Std.Dev. |  | 01: 35.78%; 02: 25.91%; 03: 37.57%; 05: 32.97%; 06: 14.41%; 07: 3.631%; 09: 4.218%
PDM3
| Z-Score (Absolute) |  | 0.27 | 0.74 | 0.2 | 0.67 | 0 | 0.59 | 2.43 | 14.23 | 1.75 | NA | NA
| Analytical results for Ce in OREAS 461 (Certified Value 3508 ppm) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 3526.9517 | 3428.7 | 3266 | 3480 | 3289.37 | 3580.4551 | 3696 | 3440 |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 3587.9826 | 3441 | 3457.7 | 3520 | 3396.74 | 3625.3739 | 3628 | 3470 |  | #N/A
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 3548.9624 | 3511.6 | 3372 | 3500 | 3440.42 | 3623.758 | 3682 | 3520 |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 3572.1857 | 3527.9 | 3364.9 | 3590 | 3406.87 | 3557.8769 | 3678.5 | 3450 |  | 3508.4834
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 3479.4713 | 3440.2 | 3435.9 | 3560 | 3372.8 | 3586.7322 | 3679.5 | 3520 |  | 76
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 3535.18 | 3436.4 | 3448.6 | 3460 | 3268.71 | 3586.461 | 3766 | 3650
| Mean | [6 empty] | 3541.789 | 3464.3 | 3390.85 | 3518.3333 | 3362.485 | 3593.4429 | 3688.3333 | 3508.3333
| Median | [6 empty] | 3542.0712 | 3440.6 | 3403.95 | 3510 | 3384.77 | 3586.5966 | 3680.75 | 3495
| Std Dev. | [6 empty] | 38.0962 | 43.4781 | 72.732 | 49.1596 | 68.5029 | 26.3287 | 44.6001 | 77.3089
| Rel.Std.Dev. | [6 empty] | 1.076% | 1.255% | 2.145% | 1.397% | 2.037% | 0.7327% | 1.209% | 2.204%
| PDM3 | [6 empty] | 0.9493% | -1.259% | -3.353% | 0.2807% | -4.161% | 2.422% | 5.126% | -0.004276%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.3 | 0.51 | 1.28 | 0.05 | 1.58 | 0.84 | 1.83 | 0.05
| Analytical results for Co in OREAS 461 (Certified Value 12.1 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 11.6 | 12 | 11.8 | 12.4 | 12 | 10.5852 | 13 | 13 | 11.7 | 13.1 | 12.0447 | 25 | 15 |  | 1
| 1 | 2 | 11.9 | 11.7 | 11.9 | 11.9 | 11.5 | 10.638 | 13 | 13 | 11.7 | 12.8 | 11.9276 | 25 | 10 |  | #N/A
| 1 | 3 | 11.9 | 12 | 12.2 | 12.3 | 12 | 10.4673 | 13.3 | 13 | 12 | 12.7 | 11.8657 | 25 | 15 |  | 16
| 1 | 4 | 11.9 | 12.2 | 11.5 | 12.1 | 11.4 | 11.3589 | 13 | 13 | 11.9 | 12.9 | 11.9513 | 25 | 15 |  | 12.1248
| 1 | 5 | 12 | 11.7 | 12.1 | 12.3 | 11.6 | 11.2319 | 12.9 | 13 | 11.8 | 12.7 | 12.126 | 25 | 10 |  | 77
| 1 | 6 | 11.8 | 12 | 12 | 12.5 | 11.8 | 11.4416 | 12.9 | 12 | 12.1 | 12.4 | 12.0241 | 25 | 15
| Mean |  | 11.85 | 11.9333 | 11.9167 | 12.25 | 11.7167 | 10.9538 | 13.0167 | 12.8333 | 11.8667 | 12.7667 | 11.9899 | 25 | 13.3333
| Median |  | 11.9 | 12 | 11.95 | 12.3 | 11.7 | 10.935 | 13 | 13 | 11.85 | 12.75 | 11.9877 | 25 | 15
| Std Dev. |  | 0.13784 | 0.19664 | 0.24833 | 0.21679 | 0.25626 | 0.43627 | 0.1472 | 0.40825 | 0.1633 | 0.23381 | 0.093297 | 0 | 2.582
| Rel.Std.Dev. |  | 1.163% | 1.648% | 2.084% | 1.77% | 2.187% | 3.983% | 1.131% | 3.181% | 1.376% | 1.831% | 0.7781% | 0% | 19.36%
| PDM3 |  | -2.266% | -1.579% | -1.717% | 1.033% | -3.366% | -9.658% | 7.356% | 5.844% | -2.129% | 5.294% | -1.113% | 106.2% | 9.967%
| Z-Score (Absolute) |  | 0.35 | 0.14 | 0.18 | 0.64 | 0.67 | 2.56 | 2.53 | 2.08 | 0.3 | 1.92 | 0 | 32.11 | 3.32
| Analytical results for Cr in OREAS 461 (Certified Value 524 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 511 | 552 | 438 | 508 | 474 | 539.6079 | 529 | 521 | 641 | 391 | 600 |  |  |  | 1
| 1 | 2 | 510 | 544 | 440 | 501 | 496 | 542.0672 | 526 | 524 | 648 | 387 | 600 |  |  |  | #N/A
| 1 | 3 | 492 | 555 | 433 | 516 | 472 | 558.2945 | 571 | 514 | 637 | 421 | 600 |  |  |  | 16
| 1 | 4 | 499 | 572 | 441 | 504 | 454 | 568.1738 | 547 | 533 | 686 | 397 | 650 |  |  |  | 523.6841
| 1 | 5 | 494 | 553 | 443 | 523 | 486 | 562.6838 | 550 | 506 | 652 | 382 | 600 |  |  |  | 78
| 1 | 6 | 504 | 552 | 441 | 514 | 456 | 568.5218 | 566 | 505 | 679 | 386 | 650
| Mean |  | 501.6667 | 554.6667 | 439.3333 | 511 | 473 | 556.5582 | 548.1667 | 517.1667 | 657.1667 | 394 | 616.6667
| Median |  | 501.5 | 552.5 | 440.5 | 511 | 473 | 560.4892 | 548.5 | 517.5 | 650 | 389 | 600
| Std Dev. |  | 8.0166 | 9.288 | 3.5024 | 8.1976 | 16.4317 | 12.7741 | 18.4544 | 10.9072 | 20.4296 | 14.1704 | 25.8199
| Rel.Std.Dev. |  | 1.598% | 1.675% | 0.7972% | 1.604% | 3.474% | 2.295% | 3.367% | 2.109% | 3.109% | 3.597% | 4.187%
| PDM3 |  | -4.204% | 5.916% | -16.11% | -2.422% | -9.678% | 6.277% | 4.675% | -1.245% | 25.49% | -24.76% | 17.76%
| Z-Score (Absolute) |  | 0.27 | 0.58 | 1.33 | 0.11 | 0.76 | 0.67 | 0.53 | 0 | 2.4 | 2.2 | 1.7
| Analytical results for Cs in OREAS 461 (Certified Value 0.77 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.51 | 0.79 | 0.81 | 0.76 | 0.73 | 0.7631 | 0.8 | 0.8 | 0.84 | 0.7 | 1.0552 | 1 | 0.7 |  | 1
| 1 | 2 | 0.72 | 0.88 | 0.85 | 0.78 | 0.71 | 0.7181 | 0.8 | 0.9 | 0.85 | 0.7 | 1.043 | 1 | 0.7 |  | #N/A
| 1 | 3 | 0.74 | 0.78 | 0.86 | 0.77 | 0.73 | 0.74091 | 0.8 | 0.9 | 0.85 | 0.74 | 0.98198 | 1 | 0.7 |  | 16
| 1 | 4 | 0.8 | 0.79 | 0.79 | 0.81 | 0.71 | 0.72047 | 0.8 | 0.9 | 0.86 | 0.72 | 0.98982 | 1 | 0.7 |  | 0.76887
| 1 | 5 | 0.64 | 0.78 | 0.85 | 0.84 | 0.72 | 0.78307 | 0.8 | 0.8 | 0.84 | 0.69 | 0.99982 | 1 | 0.7 |  | 79
| 1 | 6 | 0.55 | 0.79 | 0.82 | 0.81 | 0.71 | 0.71156 | 0.8 | 0.8 | 0.84 | 0.7 | 1.0944 | 1 | 0.7
| Mean |  | 0.66 | 0.80167 | 0.83 | 0.795 | 0.71833 | 0.73954 | 0.8 | 0.85 | 0.84667 | 0.70833 | 1.0274 | 1 | 0.7
| Median |  | 0.68 | 0.79 | 0.835 | 0.795 | 0.715 | 0.73069 | 0.8 | 0.85 | 0.845 | 0.7 | 1.0214 | 1 | 0.7
| Std Dev. |  | 0.11367 | 0.038687 | 0.027568 | 0.030166 | 0.0098319 | 0.028489 | 1.2162e-16 | 0.054772 | 0.008165 | 0.018348 | 0.044132 | 0 | 1.2162e-16
| Rel.Std.Dev. |  | 17.22% | 4.826% | 3.321% | 3.794% | 1.369% | 3.852% | 1.52e-14% | 6.444% | 0.9644% | 2.59% | 4.296% | 0% | 1.737e-14%
| PDM3 |  | -14.16% | 4.266% | 7.951% | 3.399% | -6.572% | -3.815% | 4.049% | 10.55% | 10.12% | -7.873% | 33.62% | 30.06% | -8.957%
| Z-Score (Absolute) |  | 1.28 | 0.11 | 0.43 | 0 | 0.93 | 0.67 | 0.06 | 0.67 | 0.63 | 1.13 | 2.83 | 2.49 | 1.15
| Analytical results for Cu in OREAS 461 (Certified Value 56 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 56.1 | 56.6 | 56.6 | 55 | 58 | 51.9975 | 58.3 | 69 | 57.4 | 98 | 55.2444 | 26.7 | 50 | 74 | 1
| 1 | 2 | 56.4 | 55.5 | 56.2 | 53.5 | 54.5 | 51.9646 | 58.1 | 70 | 57.2 | 96 | 56.1858 | 24.2 | 60 | 72 | #N/A
| 1 | 3 | 56 | 56.7 | 57.7 | 55.9 | 56 | 53.2252 | 58.6 | 79 | 58 | 96 | 55.8862 | 28.3 | 50 | 74 | 16
| 1 | 4 | 57.4 | 56.3 | 52.1 | 54.9 | 54.1 | 54.4716 | 58.4 | 67 | 57.3 | 100 | 55.2083 | 27.4 | 50 | 74 | 55.8138
| 1 | 5 | 56.3 | 56.1 | 56.9 | 55.4 | 56 | 53.8908 | 57.6 | 64 | 57.3 | 100 | 55.4457 | 27.8 | 60 | 74 | 80
| 1 | 6 | 56.9 | 55.6 | 56.8 | 55.6 | 55.3 | 53.5101 | 58.4 | 61 | 58.4 | 85 | 55.6601 | 27.1 | 50 | 74
| Mean |  | 56.5167 | 56.1333 | 56.05 | 55.05 | 55.65 | 53.1766 | 58.2333 | 68.3333 | 57.6 | 95.8333 | 55.6051 | 26.9167 | 53.3333 | 73.6667
| Median |  | 56.35 | 56.2 | 56.7 | 55.2 | 55.65 | 53.3677 | 58.35 | 68 | 57.35 | 97 | 55.5529 | 27.25 | 50 | 74
| Std Dev. |  | 0.53448 | 0.50067 | 1.9967 | 0.84558 | 1.3867 | 1.0159 | 0.35024 | 6.186 | 0.4858 | 5.6006 | 0.38268 | 1.4414 | 5.164 | 0.8165
| Rel.Std.Dev. |  | 0.9457% | 0.8919% | 3.562% | 1.536% | 2.492% | 1.91% | 0.6014% | 9.053% | 0.8434% | 5.844% | 0.6882% | 5.355% | 9.682% | 1.108%
| PDM3 |  | 1.259% | 0.5724% | 0.4231% | -1.369% | -0.2935% | -4.725% | 4.335% | 22.43% | 3.2% | 71.7% | -0.374% | -51.77% | -4.444% | 31.99%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 0.22 | 0.54 | 0.29 | 1.33 | 0.81 | 5.09 | 0.54 | 17.66 | 0.3 | 12.23 | 1.27 | 7.35
| Analytical results for Dy in OREAS 461 (Certified Value 33.5 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 32 | 31.7 | 34.9 | 32.4 | 34.1 | 33.6238 | 33 | 32.2 | 35.1 | 33.8 | 35.9756 | 32.9 | 34.9 |  | 1
| 1 | 2 | 34.1 | 33.1 | 34.2 | 32.7 | 32.4 | 34.4217 | 32.7 | 33.3 | 35 | 33.64 | 35.6305 | 32.9 | 31.8 |  | #N/A
| 1 | 3 | 34 | 32.6 | 34.3 | 33.5 | 33.8 | 34.1324 | 33.4 | 33.7 | 34.6 | 33.46 | 35.9598 | 33.3 | 32.8 |  | 16
| 1 | 4 | 35.4 | 32.9 | 31.1 | 33.7 | 32.4 | 34.1261 | 33.4 | 33 | 33.8 | 34.07 | 35.5468 | 33 | 32.7 |  | 33.5055
| 1 | 5 | 33.5 | 32.4 | 34.7 | 33.6 | 33.5 | 34.0991 | 32.6 | 33.9 | 34.6 | 33.34 | 35.7708 | 33.1 | 31.7 |  | 81
| 1 | 6 | 32.9 | 32.3 | 33.5 | 34.2 | 34.1 | 34.5262 | 33 | 34.1 | 34.7 | 33.24 | 36.4661 | 33.4 | 34.1
| Mean |  | 33.65 | 32.5 | 33.7833 | 33.35 | 33.3833 | 34.1549 | 33.0167 | 33.3667 | 34.6333 | 33.5917 | 35.8916 | 33.1 | 33
| Median |  | 33.75 | 32.5 | 34.25 | 33.55 | 33.65 | 34.1292 | 33 | 33.5 | 34.65 | 33.55 | 35.8653 | 33.05 | 32.75
| Std Dev. |  | 1.1572 | 0.49396 | 1.4006 | 0.67157 | 0.79352 | 0.31481 | 0.33714 | 0.69761 | 0.45898 | 0.30948 | 0.32967 | 0.20976 | 1.2712
| Rel.Std.Dev. |  | 3.439% | 1.52% | 4.146% | 2.014% | 2.377% | 0.9217% | 1.021% | 2.091% | 1.325% | 0.9213% | 0.9185% | 0.6337% | 3.852%
| PDM3 |  | 0.4311% | -3.001% | 0.8291% | -0.4642% | -0.3648% | 1.938% | -1.459% | -0.4145% | 3.366% | 0.257% | 7.121% | -1.21% | -1.509%
| Z-Score (Absolute) |  | 0.49 | 1.62 | 1.72 | 0.06 | 0 | 1.42 | 0.67 | 0.03 | 2.3 | 0.38 | 4.61 | 0.52 | 0.7
| Analytical results for Er in OREAS 461 (Certified Value 7.46 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 7.4 | 7.4 | 9.14 | 7.33 | 7.31 | 7.2592 | 7.3 | 7.3 | 11.1 | 7.14 | 12.5661 | 7.4 | 8.45 |  | 1
| 1 | 2 | 8.1 | 7.44 | 9.05 | 7.42 | 7.08 | 7.6145 | 7.2 | 7.3 | 11.4 | 7.37 | 12.4829 | 7 | 7.75 |  | #N/A
| 1 | 3 | 7.97 | 7.41 | 9.16 | 7.45 | 7.4 | 7.6086 | 7.5 | 7.4 | 11.2 | 7.57 | 12.7797 | 6.95 | 7.8 |  | 16
| 1 | 4 | 8.08 | 7.57 | 8.47 | 7.42 | 7.07 | 7.4815 | 7.6 | 7.4 | 11 | 7.43 | 12.5747 | 7.35 | 7.65 |  | 7.4597
| 1 | 5 | 7.66 | 7.52 | 9.23 | 7.54 | 7.14 | 7.2285 | 7.3 | 6.9 | 11.6 | 7.32 | 12.5352 | 7.35 | 8.65 |  | 82
| 1 | 6 | 7.7 | 7.63 | 8.99 | 7.72 | 7.38 | 7.3954 | 7.4 | 7.2 | 11.2 | 7.54 | 12.7716 | 7.4 | 8.2
| Mean |  | 7.8183 | 7.495 | 9.0067 | 7.48 | 7.23 | 7.4313 | 7.3833 | 7.25 | 11.25 | 7.395 | 12.6184 | 7.2417 | 8.0833
| Median |  | 7.835 | 7.48 | 9.095 | 7.435 | 7.225 | 7.4385 | 7.35 | 7.3 | 11.2 | 7.4 | 12.5704 | 7.35 | 8
| Std Dev. |  | 0.27745 | 0.093541 | 0.27616 | 0.1355 | 0.151 | 0.16709 | 0.1472 | 0.18708 | 0.21679 | 0.15758 | 0.12602 | 0.20837 | 0.41191
| Rel.Std.Dev. |  | 3.549% | 1.248% | 3.066% | 1.811% | 2.088% | 2.248% | 1.994% | 2.58% | 1.927% | 2.131% | 0.9987% | 2.877% | 5.096%
| PDM3 |  | 4.808% | 0.4737% | 20.74% | 0.2726% | -3.079% | -0.3802% | -1.023% | -2.811% | 50.81% | -0.8668% | 69.15% | -2.922% | 8.361%
| Z-Score (Absolute) |  | 1.37 | 0.22 | 5.96 | 0 | 0.72 | 0 | 0.17 | 0.4 | 13.53 | 0.13 | 18.37 | 0.67 | 1.48
| Analytical results for Eu in OREAS 461 (Certified Value 47.6 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 51.2 | 46.1 | 50.4 | 48.7 | 47.6 | 47.6856 | 46 | 46.9 | 44.2 | 44.11 | 53.2676 | 45.2 | 48.5 |  | 1
| 1 | 2 | 55.1 | 46.4 | 48.6 | 49 | 46.5 | 47.7181 | 46.1 | 48.6 | 45.2 | 43.9 | 53.3125 | 45.4 | 39 |  | #N/A
| 1 | 3 | 55 | 47.1 | 49.3 | 50 | 47.8 | 47.846 | 46.9 | 48.2 | 44.7 | 43.76 | 53.2948 | 45.8 | 44 |  | 16
| 1 | 4 | 55.8 | 45.6 | 45.2 | 49.7 | 46.5 | 47.6537 | 47.3 | 47.4 | 43.2 | 44.4 | 53.1393 | 45.1 | 43.4 |  | 47.6394
| 1 | 5 | 52.7 | 46 | 49.5 | 49.8 | 47.8 | 47.8968 | 46.1 | 48.3 | 43 | 43.94 | 53.127 | 46 | 44 |  | 83
| 1 | 6 | 51.4 | 44.8 | 48.2 | 51.2 | 48.4 | 47.9085 | 46.6 | 49.1 | 44.3 | 42.56 | 53.5906 | 46.3 | 45.1
| Mean |  | 53.5333 | 46 | 48.5333 | 49.7333 | 47.4333 | 47.7848 | 46.5 | 48.0833 | 44.1 | 43.7783 | 53.2886 | 45.6333 | 44
| Median |  | 53.85 | 46.05 | 48.95 | 49.75 | 47.7 | 47.782 | 46.35 | 48.25 | 44.25 | 43.92 | 53.2812 | 45.6 | 44
| Std Dev. |  | 2.0216 | 0.77201 | 1.8019 | 0.8756 | 0.77115 | 0.11232 | 0.52536 | 0.80353 | 0.85323 | 0.63594 | 0.16773 | 0.4761 | 3.0601
| Rel.Std.Dev. |  | 3.776% | 1.678% | 3.713% | 1.761% | 1.626% | 0.2351% | 1.13% | 1.671% | 1.935% | 1.453% | 0.3148% | 1.043% | 6.955%
| PDM3 |  | 12.37% | -3.441% | 1.876% | 4.395% | -0.4325% | 0.3052% | -2.392% | 0.9319% | -7.43% | -8.105% | 11.86% | -4.211% | -7.639%
| Z-Score (Absolute) |  | 2.29 | 0.54 | 0.66 | 0.86 | 0 | 0.13 | 0.35 | 0.24 | 1.25 | 1.28 | 2.19 | 0.67 | 1.29
| Analytical results for Fe in OREAS 461 (Certified Value 31.52 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 30 | 30.7 | 30.3 | 33.1 | 32.4 | 32.426 | 30.51 | 31.2 | 31.9 | 31.07 | 31.2398 | > 15 | 32.0185 | 31.9 | 1
| 1 | 2 | 29.9 | 30.5 | 30.2 | 32.6 | 32.2 | 32.4677 | 30.22 | 30.95 | 32.1 | 31.48 | 31.5378 | > 15 | 31.892 | 31.9 | #N/A
| 1 | 3 | 27.7 | 30.8 | 30.2 | 33.3 | 30.3 | 32.7117 | 30.08 | 31.52 | 32.1 | 31.14 | 31.0354 | > 15 | 31.825 | 32.3 | 16
| 1 | 4 | 28.2 | 30.8 | 30.3 | 32.9 | 29.7 | 32.8824 | 30.49 | 31.3 | 32 | 32.42 | 31.5479 | > 15 | 31.539 | 31.5 | 31.5178
| 1 | 5 | 27.7 | 30.8 | 30.1 | 33.7 | 31.5 | 32.9889 | 30.3 | 31.41 | 32.1 | 32.24 | 31.3844 | > 15 | 31.59 | 31.5 | 84
| 1 | 6 | 28.2 | 31 | 30.2 | 33.5 | 29.9 | 32.9064 | 30.4 | 31.76 | 32 | 30.96 | 30.9791 | > 15 | 31.76 | 32.8
| Mean |  | 28.6167 | 30.7667 | 30.2167 | 33.1833 | 31 | 32.7305 | 30.3333 | 31.3567 | 32.0333 | 31.5517 | 31.2874 |  | 31.7707 | 31.9833
| Median |  | 28.2 | 30.8 | 30.2 | 33.2 | 30.9 | 32.797 | 30.35 | 31.355 | 32.05 | 31.31 | 31.3121 |  | 31.7925 | 31.9
| Std Dev. |  | 1.0572 | 0.1633 | 0.075277 | 0.40208 | 1.1866 | 0.23788 | 0.16633 | 0.27761 | 0.08165 | 0.63 | 0.24527 |  | 0.18194 | 0.49967
| Rel.Std.Dev. |  | 3.694% | 0.5308% | 0.2491% | 1.212% | 3.828% | 0.7268% | 0.5484% | 0.8853% | 0.2549% | 1.997% | 0.7839% |  | 0.5727% | 1.562%
| PDM3 |  | -9.205% | -2.383% | -4.128% | 5.284% | -1.643% | 3.848% | -3.758% | -0.5113% | 1.636% | 0.1074% | -0.731% |  | 0.8025% | 1.477%
| Z-Score (Absolute) |  | 2.95 | 0.63 | 1.23 | 1.97 | 0.38 | 1.48 | 1.1 | 0 | 0.73 | 0.21 | 0.07 | NA | 0.45 | 0.67
| Analytical results for Ga in OREAS 461 (Certified Value 45.5 ppm) | [15 empty] | Certified
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 10; 08: 12; 09: 22 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS | 1
1
| 1 | 1 | 01: 47.8; 02: 47.4; 03: 47.9; 05: 60.8; 06: 35.7685; 07: 10.7; 08: 46.3011; 09: 35.4 | 1
| 1 | 2 | 01: 47.1; 02: 47.9; 03: 47.3; 05: 58.2; 06: 37.1335; 07: 10.5; 08: 46.1081; 09: 33.4 | #N/A
| 1 | 3 | 01: 48.2; 02: 48.4; 03: 47.7; 05: 60; 06: 35.9095; 07: 10.1; 08: 46.3029; 09: 34.6 | 16
| 1 | 4 | 01: 48.4; 02: 43.8; 03: 47.3; 05: 57.9; 06: 36; 07: 10.6; 08: 45.8842; 09: 33.8 | 45.5202
| 1 | 5 | 01: 47.2; 02: 48.3; 03: 48.9; 05: 59.3; 06: 34.9198; 07: 10.2; 08: 45.6738; 09: 34 | 85
| 1 | 6 | 01: 47; 02: 46.1; 03: 48.6; 05: 59.5; 06: 36.084; 07: 10.7; 08: 46.545; 09: 33.2
| Mean |  | 01: 47.6167; 02: 46.9833; 03: 47.95; 05: 59.2833; 06: 35.9692; 07: 10.4667; 08: 46.1359; 09: 34.0667
| Median |  | 01: 47.5; 02: 47.65; 03: 47.8; 05: 59.4; 06: 35.9548; 07: 10.55; 08: 46.2046; 09: 33.9
| Std Dev. |  | 01: 0.60139; 02: 1.7702; 03: 0.66858; 05: 1.0907; 06: 0.70919; 07: 0.2582; 08: 0.31631; 09: 0.8165
| Rel.Std.Dev. |  | 01: 1.263%; 02: 3.768%; 03: 1.394%; 05: 1.84%; 06: 1.972%; 07: 2.467%; 08: 0.6856%; 09: 2.397%
| PDM3 |  | 01: 4.605%; 02: 3.214%; 03: 5.338%; 05: 30.24%; 06: -20.98%; 07: -77.01%; 08: 1.352%; 09: -25.16%
| Z-Score (Absolute) |  | 01: 0.08; 02: 0.08; 03: 0.12; 05: 1.4; 06: 1.23; 07: 4.1; 08: 0.08; 09: 1.44
| Analytical results for Gd in OREAS 461 (Certified Value 98 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 88.1 | 96.9 | 127.5 | 90.9 | 102.5 | 99.1766 | 99.2 | 96.1 | 47.9 | 96.67 | 142.1213 | 96.4 | 107 |  | 1
| 1 | 2 | 97.5 | 95.3 | 117.5 | 90.7 | 97.3 | 97.0894 | 100.6 | 99.3 | 45.8 | 95.38 | 141.6797 | 95.3 | 98 |  | #N/A
| 1 | 3 | 96.9 | 95.8 | 120.5 | 93.2 | 103 | 98.2242 | 100.1 | 97.4 | 45.6 | 95.31 | 141.0163 | 97.9 | 104 |  | 16
| 1 | 4 | 100.5 | 103 | 103 | 91.8 | 97.8 | 98.8098 | 103.1 | 96.4 | 46.4 | 96.82 | 141.1854 | 97.1 | 100 |  | 97.5977
| 1 | 5 | 94.7 | 98.1 | 120.5 | 92.2 | 102 | 99.8001 | 99.8 | 100.1 | 44.8 | 94.42 | 141.7589 | 96.8 | 101 |  | 86
| 1 | 6 | 93 | 97.6 | 117 | 94.5 | 102 | 96.8614 | 99.9 | 100.4 | 46.5 | 94.04 | 142.0565 | 98.5 | 103
| Mean |  | 95.1167 | 97.7833 | 117.6667 | 92.2167 | 100.7667 | 98.3269 | 100.45 | 98.2833 | 46.1667 | 95.44 | 141.6363 | 97 | 102.1667
| Median |  | 95.8 | 97.25 | 119 | 92 | 102 | 98.517 | 100 | 98.35 | 46.1 | 95.345 | 141.7193 | 96.95 | 102
| Std Dev. |  | 4.2832 | 2.7651 | 8.1035 | 1.4414 | 2.524 | 1.1675 | 1.3751 | 1.8925 | 1.0482 | 1.1347 | 0.45089 | 1.1278 | 3.1885
| Rel.Std.Dev. |  | 4.503% | 2.828% | 6.887% | 1.563% | 2.505% | 1.187% | 1.369% | 1.926% | 2.27% | 1.189% | 0.3183% | 1.163% | 3.121%
| PDM3 |  | -2.542% | 0.1902% | 20.56% | -5.513% | 3.247% | 0.7472% | 2.923% | 0.7025% | -52.7% | -2.211% | 45.12% | -0.6124% | 4.681%
| Z-Score (Absolute) |  | 0.75 | 0.37 | 5.29 | 1.44 | 0.59 | 0.01 | 0.39 | 0 | 12.36 | 0.67 | 10.28 | 0.3 | 0.92
| Analytical results for Ge in OREAS 461 (Indicative Value 3.6 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2.35 | 4.11 | 3.61 | 2.73 | 5.3 | 0.5 | 3.8 | 6.7 | < 0.05 | 4.5 |  |  |  |  | 1
| 1 | 2 | 2.51 | 3.93 | 3.47 | 2.63 | 4.79 | 0.43227 | 3.6 | 6.3 | < 0.05 | 4.6 |  |  |  |  | 20
| 1 | 3 | 2.48 | 3.86 | 3.56 | 2.67 | 4.78 | 0.4 | 3.8 | 6.3 | < 0.05 | 4.39 |  |  |  |  | 16
| 1 | 4 | 2.83 | 3.88 | 3.13 | 2.59 | 4.99 | 0.44964 | 3.7 | 5.6 | < 0.05 | 4.42 |  |  |  |  | 3.5955
| 1 | 5 | 2.53 | 3.62 | 3.6 | 2.65 | 4.87 | 0.44734 | 3.7 | 6 | < 0.05 | 4.28 |  |  |  |  | 26
| 1 | 6 | 2.39 | 3.61 | 3.37 | 2.72 | 5.08 | 0.50329 | 4 | 7.1 | < 0.05 | 4.37
| Mean |  | 2.515 | 3.835 | 3.4567 | 2.665 | 4.9683 | 0.45542 | 3.7667 | 6.3333 |  | 4.4267
| Median |  | 2.495 | 3.87 | 3.515 | 2.66 | 4.93 | 0.44849 | 3.75 | 6.3 |  | 4.405
| Std Dev. |  | 0.16944 | 0.19191 | 0.18392 | 0.053572 | 0.19994 | 0.039961 | 0.13663 | 0.52409 |  | 0.11094
| Rel.Std.Dev. |  | 6.737% | 5.004% | 5.321% | 2.01% | 4.024% | 8.774% | 3.627% | 8.275% |  | 2.506%
| PDM3 |  | -30.05% | 6.662% | -3.86% | -25.88% | 38.18% | -87.33% | 4.762% | 76.15% |  | 23.12%
| Z-Score (Absolute) |  | 0.8 | 0.04 | 0.19 | 0.67 | 0.74 | 2.03 | 0 | 1.57 | NA | 0.4
| Analytical results for Hf in OREAS 461 (Certified Value 7.39 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 8.2 | 6.6 | 8.4 | 7.6 | 7.6 | 6.4285 | 6.2 | 7.13 | 8.92 | 6.6 |  |  |  |  | 1
| 1 | 2 | 8.7 | 6.9 | 8 | 7.3 | 7.5 | 6.2507 | 6 | 7.07 | 8.81 | 6.6 |  |  |  |  | #N/A
| 1 | 3 | 8.6 | 6.7 | 8.3 | 7.6 | 7.6 | 6.7319 | 6.2 | 7.11 | 8.85 | 7 |  |  |  |  | 16
| 1 | 4 | 8.3 | 7.2 | 7.7 | 7.7 | 7.2 | 6.9212 | 5.9 | 6.91 | 9.06 | 6.8 |  |  |  |  | 7.3912
| 1 | 5 | 8.2 | 6.9 | 8.3 | 7.9 | 7.6 | 6.8714 | 6.1 | 6.99 | 8.68 | 6.8 |  |  |  |  | 87
| 1 | 6 | 8.1 | 6.7 | 8 | 7.9 | 7.9 | 6.6261 | 6.2 | 7.25 | 8.46 | 6.8
| Mean |  | 8.35 | 6.8333 | 8.1167 | 7.6667 | 7.5667 | 6.6383 | 6.1 | 7.0767 | 8.7967 | 6.7667
| Median |  | 8.25 | 6.8 | 8.15 | 7.65 | 7.6 | 6.679 | 6.15 | 7.09 | 8.83 | 6.8
| Std Dev. |  | 0.2429 | 0.21602 | 0.26394 | 0.22509 | 0.22509 | 0.25983 | 0.12649 | 0.11776 | 0.20714 | 0.15055
| Rel.Std.Dev. |  | 2.909% | 3.161% | 3.252% | 2.936% | 2.975% | 3.914% | 2.074% | 1.664% | 2.355% | 2.225%
| PDM3 |  | 12.97% | -7.547% | 9.816% | 3.727% | 2.375% | -10.19% | -17.47% | -4.255% | 19.02% | -8.449%
| Z-Score (Absolute) |  | 1.12 | 0.53 | 0.87 | 0.38 | 0.27 | 0.74 | 1.33 | 0.27 | 1.61 | 0.6
| Analytical results for Ho in OREAS 461 (Certified Value 4.06 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.33 | 3.95 | 4.63 | 4.07 | 3.96 | 4.0917 | 4.2 | 3.9 | 5.25 | 4.02 | 5.0412 | 4.1 | 4.26 |  | 1
| 1 | 2 | 4.56 | 4.05 | 4.53 | 4.01 | 3.8 | 4.0883 | 4.2 | 4.2 | 5.33 | 4.04 | 4.9453 | 4 | 3.88 |  | #N/A
| 1 | 3 | 4.51 | 4.03 | 4.55 | 4.08 | 3.98 | 4.0844 | 4.3 | 4 | 5.19 | 4.05 | 5.0033 | 4.15 | 4 |  | 16
| 1 | 4 | 4.67 | 4.02 | 4.14 | 4.15 | 3.79 | 4.2018 | 4.2 | 3.8 | 5.04 | 4.05 | 4.991 | 4.05 | 3.94 |  | 4.0576
| 1 | 5 | 4.41 | 3.98 | 4.61 | 4.14 | 3.89 | 4.2227 | 4.1 | 4.1 | 5.01 | 4.07 | 5.0459 | 4.1 | 3.94 |  | 88
| 1 | 6 | 4.39 | 4.05 | 4.45 | 4.21 | 3.95 | 4.122 | 4.2 | 4.1 | 5.23 | 4.02 | 4.9731 | 4.1 | 4.12
| Mean |  | 4.4783 | 4.0133 | 4.485 | 4.11 | 3.895 | 4.1352 | 4.2 | 4.0167 | 5.175 | 4.0417 | 5 | 4.0833 | 4.0233
| Median |  | 4.46 | 4.025 | 4.54 | 4.11 | 3.92 | 4.1068 | 4.2 | 4.05 | 5.21 | 4.045 | 4.9971 | 4.1 | 3.97
| Std Dev. |  | 0.12561 | 0.040332 | 0.18064 | 0.070711 | 0.083126 | 0.06152 | 0.063246 | 0.1472 | 0.12518 | 0.019408 | 0.039024 | 0.05164 | 0.14166
| Rel.Std.Dev. |  | 2.805% | 1.005% | 4.028% | 1.72% | 2.134% | 1.488% | 1.506% | 3.665% | 2.419% | 0.4802% | 0.7805% | 1.265% | 3.521%
| PDM3 |  | 10.37% | -1.091% | 10.53% | 1.291% | -4.008% | 1.911% | 3.509% | -1.009% | 27.54% | -0.3929% | 23.22% | 0.6339% | -0.8448%
| Z-Score (Absolute) |  | 2.66 | 0.7 | 3.21 | 0 | 1.55 | 0.18 | 0.65 | 0.67 | 7.69 | 0.49 | 6.43 | 0.19 | 0.63
| Analytical results for In in OREAS 461 (Certified Value 0.57 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 0.557 | 0.568 | 0.615 | 0.561 | 0.586 | 0.5264 | 0.6 | 0.6 | 0.56 | 0.56 | 0.6 | 0.48 |  |  | 1
| 1 | 2 | 0.555 | 0.548 | 0.668 | 0.55 | 0.552 | 0.54801 | 0.6 | 0.58 | 0.557 | 0.54 | 0.6 | 0.52 |  |  | #N/A
| 1 | 3 | 0.555 | 0.581 | 0.693 | 0.575 | 0.574 | 0.519 | 0.6 | 0.5 | 0.545 | 0.57 | 0.6 | 0.5 |  |  | 16
| 1 | 4 | 0.567 | 0.581 | 0.656 | 0.571 | 0.576 | 0.54597 | 0.6 | 0.63 | 0.549 | 0.55 | 0.6 | 0.46 |  |  | 0.56812
| 1 | 5 | 0.565 | 0.569 | 0.676 | 0.585 | 0.588 | 0.5306 | 0.6 | 0.5 | 0.564 | 0.55 | 0.6 | 0.48 |  |  | 89
| 1 | 6 | 0.548 | 0.566 | 0.644 | 0.567 | 0.592 | 0.5343 | 0.6 | 0.6 | 0.561 | 0.53 | 0.6 | 0.48
| Mean |  | 0.55783 | 0.56883 | 0.65867 | 0.56817 | 0.578 | 0.53405 | 0.6 | 0.56833 | 0.556 | 0.55 | 0.6 | 0.48667
| Median |  | 0.556 | 0.5685 | 0.662 | 0.569 | 0.581 | 0.53245 | 0.6 | 0.59 | 0.5585 | 0.55 | 0.6 | 0.48
| Std Dev. |  | 0.0070545 | 0.012156 | 0.027186 | 0.012007 | 0.014533 | 0.011259 | 0 | 0.055287 | 0.0074297 | 0.014142 | 0 | 0.020656
| Rel.Std.Dev. |  | 1.265% | 2.137% | 4.127% | 2.113% | 2.514% | 2.108% | 0% | 9.728% | 1.336% | 2.571% | 0% | 4.244%
| PDM3 |  | -1.811% | 0.1253% | 15.94% | 0.007975% | 1.739% | -5.998% | 5.611% | 0.03731% | -2.134% | -3.19% | 5.611% | -14.34%
| Z-Score (Absolute) |  | 0.46 | 0.03 | 4 | 0 | 0.43 | 1.51 | 1.4 | 0 | 0.54 | 0.81 | 1.4 | 3.9
| Analytical results for K in OREAS 461 (Certified Value 0.233 wt.%) | [15 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.2 | 0.23 | 0.23 | 0.24 | 0.25 | 0.24 | 0.23 | 0.249 | 0.2331 | 0.23 | 0.23 | 0.23 |  |  | 1
| 1 | 2 | 0.22 | 0.29 | 0.23 | 0.23 | 0.24 | 0.23 | 0.23 | 0.245 | 0.2383 | 0.23 | 0.23 | 0.22 |  |  | #N/A
| 1 | 3 | 0.21 | 0.23 | 0.22 | 0.23 | 0.23 | 0.23 | 0.23 | 0.245 | 0.2359 | 0.24 | 0.23 | 0.22 |  |  | 16
| 1 | 4 | 0.21 | 0.24 | 0.23 | 0.24 | 0.23 | 0.24 | 0.22 | 0.24 | 0.2417 | 0.23 | 0.23 | 0.21 |  |  | 0.23312
| 1 | 5 | 0.2 | 0.23 | 0.23 | 0.24 | 0.24 | 0.24 | 0.24 | 0.245 | 0.2334 | 0.23 | 0.23 | 0.21 |  |  | 90
| 1 | 6 | 0.19 | 0.23 | 0.23 | 0.24 | 0.22 | 0.24 | 0.24 | 0.247 | 0.2343 | 0.23 | 0.23 | 0.22
| Mean |  | 0.205 | 0.24167 | 0.22833 | 0.23667 | 0.235 | 0.23667 | 0.23167 | 0.24517 | 0.23612 | 0.23167 | 0.23 | 0.21833
| Median |  | 0.205 | 0.23 | 0.23 | 0.24 | 0.235 | 0.24 | 0.23 | 0.245 | 0.2351 | 0.23 | 0.23 | 0.22
| Std Dev. |  | 0.010488 | 0.024014 | 0.0040825 | 0.005164 | 0.010488 | 0.005164 | 0.0075277 | 0.0029944 | 0.0033397 | 0.0040825 | 0 | 0.0075277
| Rel.Std.Dev. |  | 5.116% | 9.937% | 1.788% | 2.182% | 4.463% | 2.182% | 3.249% | 1.221% | 1.414% | 1.762% | 0% | 3.448%
| PDM3 |  | -12.06% | 3.664% | -2.055% | 1.52% | 0.8047% | 1.52% | -0.6252% | 5.166% | 1.284% | -0.6252% | -1.34% | -6.345%
| Z-Score (Absolute) |  | 4.65 | 0.03 | 0.61 | 0.84 | 0.55 | 0.84 | 0.03 | 2.31 | 0.74 | 0.03 | 0.32 | 2.34
| Analytical results for La in OREAS 461 (Certified Value 2476 ppm) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 2060 | 2620 | 2130 | 2200 | 2120 | 2680.0044 | 2514.3 | 2509 | 2720 | 2510.28 | 2646.4168 | 2696 | 2530 |  | 1
| 1 | 2 | 2350 | 2600 | 2140 | 2200 | 2190 | 2735.4177 | 2508.4 | 2570 | 2760 | 2529.21 | 2583.5317 | 2656 | 2530 |  | #N/A
| 1 | 3 | 2270 | 2630 | 2120 | 2220 | 2100 | 2692.6341 | 2550.9 | 2519.4 | 2720 | 2531.5 | 2696.8915 | 2702 | 2580 |  | 16
| 1 | 4 | 2330 | 2630 | 2140 | 2220 | 2010 | 2787.476 | 2539.8 | 2521.6 | 2670 | 2581.26 | 2594.2433 | 2703.5 | 2520 |  | 2476.481
| 1 | 5 | 2150 | 2640 | 2120 | 2280 | 2140 | 2801.3478 | 2527.4 | 2548.3 | 2670 | 2517.51 | 2645.7666 | 2701.5 | 2540 |  | 91
| 1 | 6 | 2080 | 2640 | 2130 | 2240 | 2010 | 2696.024 | 2508.9 | 2607.4 | 2690 | 2449.6 | 2679.0053 | 2743 | 2620
| Mean |  | 2206.6667 | 2626.6667 | 2130 | 2226.6667 | 2095 | 2732.1507 | 2524.95 | 2545.95 | 2705 | 2519.8933 | 2640.9759 | 2700.3333 | 2553.3333
| Median |  | 2210 | 2630 | 2130 | 2220 | 2110 | 2715.7209 | 2520.85 | 2534.95 | 2705 | 2523.36 | 2646.0917 | 2701.75 | 2535
| Std Dev. |  | 126.912 | 15.0555 | 8.9443 | 30.1109 | 72.3187 | 51.857 | 17.5776 | 37.4784 | 35.0714 | 42.4891 | 44.9647 | 27.6363 | 38.8158
| Rel.Std.Dev. |  | 5.751% | 0.5732% | 0.4199% | 1.352% | 3.452% | 1.898% | 0.6962% | 1.472% | 1.297% | 1.686% | 1.703% | 1.023% | 1.52%
| PDM3 |  | -10.9% | 6.064% | -13.99% | -10.09% | -15.4% | 10.32% | 1.957% | 2.805% | 9.228% | 1.753% | 6.642% | 9.039% | 3.103%
| Z-Score (Absolute) |  | 1.4 | 0.36 | 1.72 | 1.32 | 1.87 | 0.81 | 0.06 | 0.03 | 0.69 | 0.08 | 0.42 | 0.67 | 0
| Analytical results for Li in OREAS 461 (Certified Value 12 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 1
1
| 1 | 1 | 9.2 | 11.9 | 10.4 | 12.6 | 11.2 | 12.9489 | 12.1 | 13 | 11.8 | 10.9 | 14.2544 | 13 | 10 | 12 | 1
| 1 | 2 | 9.9 | 12.5 | 10.5 | 12.3 | 10.6 | 12.6566 | 12.1 | 13 | 11.8 | 10.4 | 14.187 | 12 | 15 | 11.5 | #N/A
| 1 | 3 | 10 | 12.2 | 11.1 | 12.6 | 11 | 13.3867 | 12.1 | 13 | 12 | 10 | 13.8567 | 13 | 10 | 12 | 16
| 1 | 4 | 10.1 | 11.8 | 12.3 | 12.7 | 10 | 13.0297 | 12.1 | 13 | 11.9 | 10.5 | 13.3227 | 14 | 15 | 11.5 | 12.0397
| 1 | 5 | 10 | 11.7 | 10.9 | 13 | 10.4 | 13.6092 | 11.8 | 14 | 11.6 | 10.5 | 13.5922 | 13 | 15 | 12 | 92
| 1 | 6 | 9.9 | 11.8 | 10.8 | 13 | 10 | 13.2865 | 11.9 | 13 | 12 | 10.4 | 14.1885 | 13 | 15 | 11.5
| Mean |  | 9.85 | 11.9833 | 11 | 12.7 | 10.5333 | 13.1529 | 12.0167 | 13.1667 | 11.85 | 10.45 | 13.9002 | 13 | 13.3333 | 11.75
| Median |  | 9.95 | 11.85 | 10.85 | 12.65 | 10.5 | 13.1581 | 12.1 | 13 | 11.85 | 10.45 | 14.0219 | 13 | 15 | 11.75
| Std Dev. |  | 0.32711 | 0.30605 | 0.68702 | 0.26833 | 0.50067 | 0.3417 | 0.13292 | 0.40825 | 0.15166 | 0.2881 | 0.37978 | 0.63246 | 2.582 | 0.27386
| Rel.Std.Dev. |  | 3.321% | 2.554% | 6.246% | 2.113% | 4.753% | 2.598% | 1.106% | 3.101% | 1.28% | 2.757% | 2.732% | 4.865% | 19.36% | 2.331%
| PDM3 |  | -18.19% | -0.4686% | -8.636% | 5.484% | -12.51% | 9.246% | -0.1917% | 9.36% | -1.576% | -13.2% | 15.45% | 7.976% | 10.74% | -2.407%
| Z-Score (Absolute) |  | 1.17 | 0.01 | 0.73 | 0.41 | 0.85 | 0.67 | 0.01 | 0.68 | 0.09 | 0.9 | 1.1 | 0.58 | 0.78 | 0.15
| Analytical results for Lu in OREAS 461 (Certified Value 0.39 ppm) | [15 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 0.52 | 0.4 | 0.41 | 0.34 | 0.39 | 0.36128 | 0.3 | 0.38 | 0.452 | 0.34 | 0.44422 | 0.3 | 0.42 |  | 1
| 1 | 2 | 0.56 | 0.39 | 0.41 | 0.33 | 0.38 | 0.352 | 0.3 | 0.29 | 0.459 | 0.33 | 0.37692 | 0.3 | 0.4 |  | #N/A
| 1 | 3 | 0.55 | 0.39 | 0.42 | 0.35 | 0.4 | 0.35677 | 0.3 | 0.3 | 0.466 | 0.34 | 0.45734 | 0.3 | 0.42 |  | 16
| 1 | 4 | 0.54 | 0.42 | 0.39 | 0.36 | 0.39 | 0.36321 | 0.3 | 0.29 | 0.454 | 0.33 | 0.41741 | 0.4 | 0.42 |  | 0.38585
| 1 | 5 | 0.51 | 0.4 | 0.43 | 0.35 | 0.4 | 0.36308 | 0.3 | 0.28 | 0.466 | 0.38 | 0.40154 | 0.3 | 0.42 |  | 93
| 1 | 6 | 0.51 | 0.39 | 0.4 | 0.37 | 0.4 | 0.37721 | 0.4 | 0.32 | 0.473 | 0.33 | 0.40907 | 0.4 | 0.44
| Mean |  | 0.53167 | 0.39833 | 0.41 | 0.35 | 0.39333 | 0.36226 | 0.31667 | 0.31 | 0.46167 | 0.34167 | 0.41775 | 0.33333 | 0.42
| Median |  | 0.53 | 0.395 | 0.41 | 0.35 | 0.395 | 0.36218 | 0.3 | 0.295 | 0.4625 | 0.335 | 0.41324 | 0.3 | 0.42
| Std Dev. |  | 0.02137 | 0.01169 | 0.014142 | 0.014142 | 0.008165 | 0.008502 | 0.040825 | 0.036878 | 0.0080664 | 0.019408 | 0.029235 | 0.05164 | 0.012649
| Rel.Std.Dev. |  | 4.019% | 2.935% | 3.449% | 4.041% | 2.076% | 2.347% | 12.89% | 11.9% | 1.747% | 5.68% | 6.998% | 15.49% | 3.012%
| PDM3 |  | 37.79% | 3.236% | 6.26% | -9.29% | 1.94% | -6.113% | -17.93% | -19.66% | 19.65% | -11.45% | 8.268% | -13.61% | 8.851%
| Z-Score (Absolute) |  | 2.15 | 0.08 | 0.26 | 0.67 | 0 | 0.48 | 1.45 | 1.51 | 1.06 | 0.8 | 0.38 | 0.93 | 0.41
| Analytical results for Mg in OREAS 461 (Certified Value 0.999 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.89 | 0.94 | 0.83 | 0.88 | 0.83 | 1.1421 | 1.01 | 1.04 | 1.08 | 1.1136 | 1.069 | 0.91 | 1.082 | 1.09 | 1
| 1 | 2 | 0.97 | 0.93 | 0.83 | 0.87 | 0.85 | 1.1798 | 1.01 | 1.04 | 1.09 | 1.1414 | 1.0852 | 0.92 | 1.0815 | 1.1 | #N/A
| 1 | 3 | 0.93 | 0.95 | 0.83 | 0.88 | 0.82 | 1.1297 | 0.99 | 1.05 | 1.09 | 1.1226 | 1.0593 | 0.92 | 1.0785 | 1.12 | 16
| 1 | 4 | 0.95 | 0.95 | 0.83 | 0.87 | 0.78 | 1.135 | 1.01 | 1.03 | 1.09 | 1.1684 | 1.0515 | 0.94 | 1.075 | 1.1 | 0.99901
| 1 | 5 | 0.9 | 0.95 | 0.83 | 0.91 | 0.84 | 1.1427 | 1.02 | 1.04 | 1.09 | 1.1137 | 1.052 | 0.93 | 1.0825 | 1.11 | 94
| 1 | 6 | 0.89 | 0.95 | 0.83 | 0.9 | 0.78 | 1.1812 | 1.01 | 1.06 | 1.08 | 1.118 | 1.0535 | 0.95 | 1.0755 | 1.12
| Mean |  | 0.92167 | 0.945 | 0.83 | 0.885 | 0.81667 | 1.1518 | 1.0083 | 1.0433 | 1.0867 | 1.1296 | 1.0617 | 0.92833 | 1.0792 | 1.1067
| Median |  | 0.915 | 0.95 | 0.83 | 0.88 | 0.825 | 1.1424 | 1.01 | 1.04 | 1.09 | 1.1203 | 1.0564 | 0.925 | 1.08 | 1.105
| Std Dev. |  | 0.033714 | 0.0083666 | 0 | 0.016432 | 0.030111 | 0.022785 | 0.0098319 | 0.010328 | 0.005164 | 0.02162 | 0.013236 | 0.01472 | 0.0033417 | 0.012111
| Rel.Std.Dev. |  | 3.658% | 0.8854% | 0% | 1.857% | 3.687% | 1.978% | 0.9751% | 0.9899% | 0.4752% | 1.914% | 1.247% | 1.586% | 0.3097% | 1.094%
| PDM3 |  | -7.742% | -5.407% | -16.92% | -11.41% | -18.25% | 15.29% | 0.9329% | 4.436% | 8.774% | 13.07% | 6.279% | -7.075% | 8.023% | 10.78%
| Z-Score (Absolute) |  | 0.79 | 0.62 | 1.49 | 1.07 | 1.59 | 0.96 | 0.13 | 0.13 | 0.46 | 0.73 | 0.27 | 0.74 | 0.41 | 0.62
| Analytical results for Mn in OREAS 461 (Certified Value 547 ppm) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 500 | 543 | 421 | 422 | 435 | 580.4329 | 632.7 | 530 | 548 | 629 | 835 | 580 | 670 |  | 1
| 1 | 2 | 546 | 551 | 421 | 422 | 466 | 579.6071 | 630.8 | 530 | 552 | 650 | 838 | 580 | 664 |  | #N/A
| 1 | 3 | 515 | 549 | 412 | 441 | 461 | 577.9823 | 634.2 | 547 | 554 | 633 | 843 | 580 | 686 |  | 16
| 1 | 4 | 517 | 572 | 418 | 440 | 434 | 596.4188 | 627.6 | 539 | 553 | 655 | 861 | 575 | 664 |  | 547.2361
| 1 | 5 | 517 | 550 | 423 | 466 | 445 | 599.495 | 617.8 | 543 | 566 | 625 | 820 | 580 | 658 |  | 95
| 1 | 6 | 517 | 546 | 425 | 463 | 412 | 597.4664 | 625.5 | 553 | 562 | 628 | 845 | 590 | 684
| Mean |  | 518.6667 | 551.8333 | 420 | 442.3333 | 442.1667 | 588.5671 | 628.1 | 540.3333 | 555.8333 | 636.6667 | 840.3333 | 580.8333 | 671
| Median |  | 517 | 549.5 | 421 | 440.5 | 440 | 588.4258 | 629.2 | 541 | 553.5 | 631 | 840.5 | 580 | 667
| Std Dev. |  | 14.9488 | 10.3037 | 4.5607 | 19.0858 | 19.7931 | 10.1858 | 5.9793 | 9.2448 | 6.7651 | 12.628 | 13.4412 | 4.916 | 11.5065
| Rel.Std.Dev. |  | 2.882% | 1.867% | 1.086% | 4.315% | 4.476% | 1.731% | 0.952% | 1.711% | 1.217% | 1.983% | 1.6% | 0.8464% | 1.715%
| PDM3 |  | -5.221% | 0.8401% | -23.25% | -19.17% | -19.2% | 7.553% | 14.78% | -1.261% | 1.571% | 16.34% | 53.56% | 6.139% | 22.62%
| Z-Score (Absolute) |  | 0.4 | 0.07 | 1.27 | 1.06 | 1.06 | 0.31 | 0.67 | 0.14 | 0 | 0.75 | 2.65 | 0.23 | 1.07
| Analytical results for Mo in OREAS 461 (Certified Value 48.2 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 47.1 | 47.8 | 53.9 | 48.5 | 50 | 44.3617 | 46.2 | 50 | 48.6 | 49.5 | 44.3514 | 42 |  |  | 1
| 1 | 2 | 46.4 | 49.3 | 53.5 | 48.1 | 48.4 | 43.7558 | 46.7 | 49 | 48.5 | 49.9 | 44.4944 | 41.5 |  |  | #N/A
| 1 | 3 | 46.1 | 48.9 | 55.4 | 48.7 | 51.2 | 45.1128 | 47.3 | 48 | 48.8 | 50.2 | 44.8468 | 42 |  |  | 16
| 1 | 4 | 47.2 | 48.5 | 50.5 | 49.3 | 49.5 | 45.9679 | 46.8 | 50 | 49 | 49.8 | 44.8924 | 41.5 |  |  | 48.2464
| 1 | 5 | 46 | 47.6 | 54.8 | 50.1 | 49.4 | 45.3298 | 46.5 | 47 | 47.8 | 49.1 | 44.159 | 42 |  |  | 96
| 1 | 6 | 47.1 | 47.5 | 52.9 | 49.6 | 51.1 | 44.8604 | 46.7 | 49 | 48.8 | 49.6 | 44.9333 | 42
| Mean |  | 46.65 | 48.2667 | 53.5 | 49.05 | 49.9333 | 44.8981 | 46.7 | 48.8333 | 48.5833 | 49.6833 | 44.6129 | 41.8333
| Median |  | 46.75 | 48.15 | 53.7 | 49 | 49.75 | 44.9866 | 46.7 | 49 | 48.7 | 49.7 | 44.6706 | 42
| Std Dev. |  | 0.54681 | 0.74476 | 1.7216 | 0.74766 | 1.0764 | 0.77082 | 0.36332 | 1.169 | 0.42151 | 0.37639 | 0.3237 | 0.2582
| Rel.Std.Dev. |  | 1.172% | 1.543% | 3.218% | 1.524% | 2.156% | 1.717% | 0.778% | 2.394% | 0.8676% | 0.7576% | 0.7256% | 0.6172%
| PDM3 |  | -3.309% | 0.04191% | 10.89% | 1.666% | 3.496% | -6.94% | -3.205% | 1.216% | 0.6983% | 2.978% | -7.531% | -13.29%
| Z-Score (Absolute) |  | 0.74 | 0.07 | 2.12 | 0.26 | 0.63 | 1.47 | 0.72 | 0.17 | 0.07 | 0.52 | 1.59 | 2.75
| Analytical results for Na in OREAS 461 (Certified Value 0.178 wt.%) | [15 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.17 | 0.16 | 0.17 | 0.19 | 0.18 | 0.18 | 0.17 | 0.184 | 0.183 | 0.20266 | 0.17 | 0.17 | 0.2 |  | 1
| 1 | 2 | 0.17 | 0.36 | 0.17 | 0.19 | 0.18 | 0.17 | 0.17 | 0.184 | 0.1859 | 0.20186 | 0.17 | 0.17 | 0.2 |  | 16
| 1 | 3 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.17 | 0.18 | 0.184 | 0.1858 | 0.20449 | 0.17 | 0.17 | 0.2 |  | 16
| 1 | 4 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.17 | 0.181 | 0.1918 | 0.20443 | 0.18 | 0.17 | 0.19 |  | 0.17817
| 1 | 5 | 0.15 | 0.16 | 0.17 | 0.19 | 0.18 | 0.18 | 0.17 | 0.184 | 0.1826 | 0.20465 | 0.17 | 0.17 | 0.19 |  | 97
| 1 | 6 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.17 | 0.186 | 0.1847 | 0.20274 | 0.17 | 0.17 | 0.2
| Mean |  | 0.16167 | 0.19333 | 0.17 | 0.19 | 0.175 | 0.17667 | 0.17167 | 0.18383 | 0.18563 | 0.20347 | 0.17167 | 0.17 | 0.19667
| Median |  | 0.16 | 0.16 | 0.17 | 0.19 | 0.175 | 0.18 | 0.17 | 0.184 | 0.18525 | 0.20358 | 0.17 | 0.17 | 0.2
| Std Dev. |  | 0.0075277 | 0.08165 | 0 | 3.0405e-17 | 0.0054772 | 0.005164 | 0.0040825 | 0.0016021 | 0.0033206 | 0.001195 | 0.0040825 | 3.0405e-17 | 0.005164
| Rel.Std.Dev. |  | 4.656% | 42.23% | 0% | 1.6e-14% | 3.13% | 2.923% | 2.378% | 0.8715% | 1.789% | 0.5873% | 2.378% | 1.789e-14% | 2.626%
| PDM3 |  | -9.265% | 8.508% | -4.588% | 6.637% | -1.782% | -0.8464% | -3.653% | 3.176% | 4.186% | 14.2% | -3.653% | -4.588% | 10.38%
| Z-Score (Absolute) |  | 1.02 | 1.15 | 0.38 | 1.15 | 0 | 0.13 | 0.25 | 0.67 | 0.81 | 2.17 | 0.25 | 0.38 | 1.65
| Analytical results for Nb in OREAS 461 (Indicative Value 1153 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 1219.5947 | 1186.1 | 906 | > 500 | 1290 |  |  |  |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 1209.887 | 1177.1 | 871 | > 500 | 1280 |  |  |  |  | 20
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 1237 | 1184.5 | 930 | > 500 | 1360 |  |  |  |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 1244.61 | 1150.5 | 920 | > 500 | 1320 |  |  |  |  | 1152.7857
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 1268.403 | 1123.2 | 912 | > 500 | 1330 |  |  |  |  | 26
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 1238.4612 | 1140.5 | 908 | > 500 | 1260
| Mean |  | 07: 1236.326; 08: 1160.3167; 09: 907.8333; 22: 1306.6667
| Median |  | 07: 1237.7306; 08: 1163.8; 09: 910; 22: 1305
| Std Dev. |  | 07: 20.4036; 08: 26.0692; 09: 20.084; 22: 36.697
| Rel.Std.Dev. |  | 07: 1.65%; 08: 2.247%; 09: 2.212%; 22: 2.808%
| PDM3 |  | 07: 7.247%; 08: 0.6533%; 09: -21.25%; 22: 13.35%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.35 | 0.35 | 2.68 | NA | 1
| Analytical results for Nd in OREAS 461 (Certified Value 1623 ppm) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 1505 | 1720 | 1520 | 1650 | 1550 | 1749.267 | 1597.8 | 1546.2 | 1760 | 1593.22 | 1523.459 | 1675 | 1710 |  | 1
| 1 | 2 | 1625 | 1645 | 1580 | 1635 | 1495 | 1707.1995 | 1561.3 | 1612.6 | 1770 | 1552.38 | 1494.7687 | 1661 | 1590 |  | #N/A
| 1 | 3 | 1610 | 1710 | 1575 | 1680 | 1550 | 1710.5504 | 1615.7 | 1592.5 | 1730 | 1595.91 | 1563.0645 | 1670.5 | 1610 |  | 16
| 1 | 4 | 1675 | 1720 | 1540 | 1755 | 1505 | 1748.9323 | 1620.2 | 1587.8 | 1720 | 1584.81 | 1485.6337 | 1670.5 | 1570 |  | 1623.483
| 1 | 5 | 1555 | 1680 | 1605 | 1775 | 1525 | 1664.5483 | 1587.2 | 1618 | 1700 | 1555.3 | 1526.8783 | 1655.5 | 1570 |  | 98
| 1 | 6 | 1505 | 1690 | 1570 | 1725 | 1555 | 1741.3026 | 1593.8 | 1641.2 | 1750 | 1504.48 | 1550.6707 | 1692.5 | 1670
| Mean |  | 1579.1667 | 1694.1667 | 1565 | 1703.3333 | 1530 | 1720.3 | 1596 | 1599.7167 | 1738.3333 | 1564.35 | 1524.0791 | 1670.8333 | 1620
| Median |  | 1582.5 | 1700 | 1572.5 | 1702.5 | 1537.5 | 1725.9265 | 1595.8 | 1602.55 | 1740 | 1570.055 | 1525.1686 | 1670.5 | 1600
| Std Dev. |  | 69.0229 | 29.0545 | 30.3315 | 57.1548 | 25.6905 | 33.0906 | 21.276 | 32.4938 | 26.3944 | 34.8072 | 30.2402 | 12.7971 | 57.6194
| Rel.Std.Dev. |  | 4.371% | 1.715% | 1.938% | 3.355% | 1.679% | 1.924% | 1.333% | 2.031% | 1.518% | 2.225% | 1.984% | 0.7659% | 3.557%
| PDM3 |  | -2.73% | 4.354% | -3.602% | 4.918% | -5.758% | 5.964% | -1.693% | -1.464% | 7.074% | -3.642% | -6.123% | 2.917% | -0.2145%
| Z-Score (Absolute) |  | 0.2 | 0.91 | 0.34 | 1 | 0.67 | 1.17 | 0.04 | 0 | 1.34 | 0.34 | 0.73 | 0.69 | 0.2
| Analytical results for Ni in OREAS 461 (Certified Value 71 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
1
| 1 | 1 | 72 | 68.8 | 70.9 | 71.3 | 69.3 | 65.2084 | 72 | 72 | 77 | 70 | 86 |  |  |  | 1
| 1 | 2 | 72 | 67.2 | 72.8 | 68.9 | 67.3 | 67.5234 | 74 | 72.6 | 73 | 70 | 86 |  |  |  | #N/A
| 1 | 3 | 72 | 68.9 | 73.1 | 71.9 | 69.2 | 66.2823 | 77 | 71.3 | 75 | 70 | 94 |  |  |  | 16
| 1 | 4 | 73 | 69.3 | 69 | 69.1 | 65.2 | 69.8188 | 82 | 71.3 | 79 | 70 | 94 |  |  |  | 71.1258
| 1 | 5 | 72.5 | 67.1 | 72.3 | 72.1 | 66.7 | 69.0167 | 88 | 70.9 | 75 | 70 | 86 |  |  |  | 99
| 1 | 6 | 72 | 69 | 72.4 | 72.2 | 66.8 | 65.5968 | 75 | 72.4 | 75 | 70 | 86
| Mean |  | 72.25 | 68.3833 | 71.75 | 70.9167 | 67.4167 | 67.2411 | 78 | 71.75 | 75.6667 | 70 | 88.6667
| Median |  | 72 | 68.85 | 72.35 | 71.6 | 67.05 | 66.9029 | 76 | 71.65 | 75 | 70 | 86
| Std Dev. |  | 0.41833 | 0.9704 | 1.545 | 1.5184 | 1.5842 | 1.8779 | 5.9666 | 0.68337 | 2.0656 | 0 | 4.1312
| Rel.Std.Dev. |  | 0.579% | 1.419% | 2.153% | 2.141% | 2.35% | 2.793% | 7.649% | 0.9524% | 2.73% | 0% | 4.659%
| PDM3 |  | 1.581% | -3.856% | 0.8776% | -0.294% | -5.215% | -5.462% | 9.665% | 0.8776% | 6.384% | -1.583% | 24.66%
| Z-Score (Absolute) |  | 0.1 | 0.7 | 0.11 | 0.17 | 0.9 | 0.94 | 0.88 | 0 | 0.67 | 0.36 | 3.51
| Analytical results for P in OREAS 461 (Certified Value 0.338 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.333 | 0.345 | 0.3 | 0.324 | 0.312 | 0.3411 | 0.37 | 0.337 | 0.3628 | 0.39836 | 0.3274 | 0.3635 | 0.335 |  | 1
| 1 | 2 | 0.336 | 0.344 | 0.301 | 0.322 | 0.321 | 0.3409 | 0.37 | 0.336 | 0.3622 | 0.40359 | 0.3123 | 0.365 | 0.345 |  | #N/A
| 1 | 3 | 0.322 | 0.348 | 0.301 | 0.329 | 0.307 | 0.3376 | 0.37 | 0.338 | 0.3624 | 0.39544 | 0.3253 | 0.3635 | 0.355 |  | 16
| 1 | 4 | 0.328 | 0.348 | 0.301 | 0.326 | 0.294 | 0.3431 | 0.37 | 0.33 | 0.3722 | 0.39672 | 0.3238 | 0.363 | 0.345 |  | 0.33785
| 1 | 5 | 0.32 | 0.349 | 0.3 | 0.335 | 0.316 | 0.3429 | 0.37 | 0.335 | 0.357 | 0.40116 | 0.3217 | 0.36 | 0.355 |  | 100
| 1 | 6 | 0.324 | 0.35 | 0.3 | 0.332 | 0.294 | 0.3433 | 0.38 | 0.337 | 0.35 | 0.39796 | 0.3234 | 0.364 | 0.345
| Mean |  | 0.32717 | 0.34733 | 0.3005 | 0.328 | 0.30733 | 0.34148 | 0.37167 | 0.3355 | 0.3611 | 0.39887 | 0.32232 | 0.36317 | 0.34667
| Median |  | 0.326 | 0.348 | 0.3005 | 0.3275 | 0.3095 | 0.342 | 0.37 | 0.3365 | 0.3623 | 0.39816 | 0.3236 | 0.3635 | 0.345
| Std Dev. |  | 0.0063377 | 0.0023381 | 0.00054772 | 0.0049396 | 0.011308 | 0.0021674 | 0.0040825 | 0.002881 | 0.0073351 | 0.0029997 | 0.0052693 | 0.0016931 | 0.0075277
| Rel.Std.Dev. |  | 1.937% | 0.6732% | 0.1823% | 1.506% | 3.679% | 0.6347% | 1.098% | 0.8587% | 2.031% | 0.752% | 1.635% | 0.4662% | 2.171%
| PDM3 |  | -3.163% | 2.806% | -11.06% | -2.916% | -9.033% | 1.075% | 10.01% | -0.6965% | 6.881% | 18.06% | -4.599% | 7.492% | 2.609%
| Z-Score (Absolute) |  | 0.56 | 0.23 | 1.61 | 0.53 | 1.34 | 0 | 1.19 | 0.24 | 0.77 | 2.25 | 0.67 | 0.85 | 0.2
| Analytical results for Pb in OREAS 461 (Certified Value 104 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 101.5 | 101 | 113 | 104.5 | 108.5 | 104.4025 | 101 | 108 | 90.9 | 173.189 | 140 | 97 |  |  | 1
| 1 | 2 | 102 | 105 | 113.5 | 104 | 106 | 106.0689 | 102 | 108 | 92.4 | 172.839 | 140 | 100 |  |  | #N/A
| 1 | 3 | 98.4 | 103 | 114 | 109 | 107.5 | 105.7455 | 104 | 108 | 95 | 169.819 | 140 | 103 |  |  | 16
| 1 | 4 | 101 | 107 | 112.5 | 107.5 | 104.5 | 107.7794 | 103 | 105 | 91.5 | 170.853 | 130 | 102 |  |  | 103.9184
| 1 | 5 | 96.7 | 103.5 | 117 | 109 | 106 | 107.1044 | 102 | 107 | 106.7 | 169.541 | 150 | 100 |  |  | 101
| 1 | 6 | 99.6 | 103 | 113 | 108 | 106 | 105.4828 | 104 | 108 | 87.8 | 171.023 | 140 | 106
| Mean |  | 99.8667 | 103.75 | 113.8333 | 107 | 106.4167 | 106.0972 | 102.6667 | 107.3333 | 94.05 | 171.2107 | 140 | 101.3333
| Median |  | 100.3 | 103.25 | 113.25 | 107.75 | 106 | 105.9072 | 102.5 | 108 | 91.95 | 170.938 | 140 | 101
| Std Dev. |  | 2.0373 | 2.0433 | 1.633 | 2.2136 | 1.3934 | 1.2015 | 1.2111 | 1.2111 | 6.6202 | 1.5135 | 6.3246 | 3.0768
| Rel.Std.Dev. |  | 2.04% | 1.969% | 1.435% | 2.069% | 1.309% | 1.132% | 1.18% | 1.128% | 7.039% | 0.884% | 4.518% | 3.036%
| PDM3 |  | -3.899% | -0.162% | 9.541% | 2.965% | 2.404% | 2.097% | -1.205% | 3.286% | -9.496% | 64.75% | 34.72% | -2.488%
| Z-Score (Absolute) |  | 1.01 | 0.4 | 1.1 | 0.12 | 0.03 | 0.03 | 0.57 | 0.17 | 2.33 | 10.29 | 5.34 | 0.78
| Analytical results for Pr in OREAS 461 (Certified Value 482 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 472 | 463 | 464 | 478 | 461 | 507.7527 | 466.4 | 465.68 | 502 | 473.31 | 493.3063 | 493.1 | 496 |  | 1
| 1 | 2 | 507 | 461 | 484 | 474 | 436 | 508.1749 | 465.8 | 482.07 | 507 | 477.84 | 492.056 | 480.8 | 454 |  | #N/A
| 1 | 3 | 507 | 469 | 490 | 483 | 455 | 506.6946 | 477.7 | 473.98 | 507 | 489.77 | 491.9507 | 488.3 | 463 |  | 16
| 1 | 4 | 528 | 467 | 472 | 509 | 439 | 522.013 | 476.8 | 474.55 | 501 | 490.32 | 490.0218 | 489.7 | 454 |  | 482.2815
| 1 | 5 | 488 | 456 | 496 | 515 | 446 | 519.6 | 474.9 | 484.41 | 516 | 475.92 | 493.2285 | 488.7 | 459 |  | 102
| 1 | 6 | 471 | 451 | 482 | 503 | 458 | 505.5371 | 473.1 | 488.12 | 501 | 466.57 | 490.684 | 496.5 | 473
| Mean |  | 495.5 | 461.1667 | 481.3333 | 493.6667 | 449.1667 | 511.6287 | 472.45 | 478.135 | 505.6667 | 478.955 | 491.8746 | 489.5167 | 466.5
| Median |  | 497.5 | 462 | 483 | 493 | 450.5 | 507.9638 | 474 | 478.31 | 504.5 | 476.88 | 492.0033 | 489.2 | 461
| Std Dev. |  | 22.4922 | 6.7651 | 11.7075 | 17.4547 | 10.3811 | 7.208 | 5.1718 | 8.2438 | 5.785 | 9.4004 | 1.3245 | 5.2871 | 16.0842
| Rel.Std.Dev. |  | 4.539% | 1.467% | 2.432% | 3.536% | 2.311% | 1.409% | 1.095% | 1.724% | 1.144% | 1.963% | 0.2693% | 1.08% | 3.448%
| PDM3 |  | 2.741% | -4.378% | -0.1966% | 2.361% | -6.866% | 6.085% | -2.039% | -0.8598% | 4.849% | -0.6897% | 1.989% | 1.5% | -3.272%
| Z-Score (Absolute) |  | 0.77 | 1.1 | 0 | 0.67 | 1.76 | 1.66 | 0.49 | 0.17 | 1.33 | 0.13 | 0.58 | 0.45 | 1.13
| Analytical results for Rb in OREAS 461 (Certified Value 13.3 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 7.2 | 14.2 | 14.3 | 13.1 | 13.6 | 12.7533 | 13.2 | 15.2 | 11.4 | 14.36 | 17.7266 | 13 | 14.4 |  | 1
| 1 | 2 | 12.4 | 13.9 | 14.3 | 12.9 | 12.9 | 12.3943 | 13 | 14.2 | 11.3 | 14.14 | 17.2784 | 13 | 13.6 |  | #N/A
| 1 | 3 | 13.2 | 14.1 | 14.6 | 13 | 13.5 | 12.6891 | 12.9 | 14.7 | 11.5 | 14.43 | 17.2067 | 13 | 13.8 |  | 16
| 1 | 4 | 15.6 | 14.3 | 13.7 | 13.3 | 13 | 12.9409 | 13.1 | 14 | 11.4 | 14.35 | 17.5604 | 13 | 13.6 |  | 13.2859
| 1 | 5 | 9.7 | 13.9 | 14.4 | 13.4 | 13.3 | 12.6223 | 13 | 13.9 | 11.3 | 14 | 17.8282 | 12.5 | 13.6 |  | 103
| 1 | 6 | 8.4 | 14.1 | 13.8 | 13.4 | 13.4 | 12.2251 | 13.1 | 14.1 | 11.4 | 13.65 | 17.8275 | 13 | 13.4
| Mean |  | 11.0833 | 14.0833 | 14.1833 | 13.1833 | 13.2833 | 12.6042 | 13.05 | 14.35 | 11.3833 | 14.155 | 17.5713 | 12.9167 | 13.7333
| Median |  | 11.05 | 14.1 | 14.3 | 13.2 | 13.35 | 12.6557 | 13.05 | 14.15 | 11.4 | 14.245 | 17.6435 | 13 | 13.6
| Std Dev. |  | 3.1877 | 0.16021 | 0.35449 | 0.2137 | 0.27869 | 0.25726 | 0.10488 | 0.501 | 0.075277 | 0.29494 | 0.27371 | 0.20412 | 0.35024
| Rel.Std.Dev. |  | 28.76% | 1.138% | 2.499% | 1.621% | 2.098% | 2.041% | 0.8037% | 3.491% | 0.6613% | 2.084% | 1.558% | 1.58% | 2.55%
| PDM3 |  | -16.58% | 6.002% | 6.755% | -0.772% | -0.01935% | -5.131% | -1.776% | 8.009% | -14.32% | 6.541% | 32.26% | -2.779% | 3.368%
| Z-Score (Absolute) |  | 0.55 | 0.79 | 0.89 | 0.1 | 0 | 0.67 | 0.23 | 0.89 | 1.89 | 0.87 | 4.26 | 0.28 | 0.31
| Analytical results for Re in OREAS 461 (Certified Value < 0.01 ppm) | [15 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.008 | < 0.002 | 0.004 | < 0.002 | 0.003 | < 0.002 | < 0.02 | < 0.002 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.01 | < 0.002 | 0.005 | < 0.002 | 0.003 | < 0.002 | < 0.02 | 0.002 | 0.003 | < 0.05 | < 0.1 |  |  |  | 17
| 1 | 3 | 0.008 | 0.002 | 0.004 | 0.002 | 0.003 | < 0.002 | < 0.02 | < 0.002 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.008 | 0.003 | 0.003 | 0.005 | 0.003 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1 |  |  |  | < 0.01
| 1 | 5 | 0.007 | 0.006 | 0.004 | 0.003 | 0.004 | < 0.002 | < 0.02 | 0.002 | 0.004 | < 0.05 | < 0.1 |  |  |  | 104
| 1 | 6 | 0.007 | < 0.002 | 0.003 | 0.002 | 0.003 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1
| Mean |  | 01: 0.008; 02: 0.0036667; 03: 0.0038333; 05: 0.003; 06: 0.0031667; 09: 0.0025; 10: 0.0035
| Median |  | 01: 0.008; 02: 0.003; 03: 0.004; 05: 0.0025; 06: 0.003; 09: 0.0025; 10: 0.0035
| Std Dev. |  | 01: 0.0010954; 02: 0.0020817; 03: 0.00075277; 05: 0.0014142; 06: 0.00040825; 09: 0.00057735; 10: 0.00070711
| Rel.Std.Dev. |  | 01: 13.69%; 02: 56.77%; 03: 19.64%; 05: 47.14%; 06: 12.89%; 09: 23.09%; 10: 20.2%
PDM3
| Z-Score (Absolute) |  | 4.59 | 0.59 | 1.21 | 0 | 0.67 | NA | NA | 0.27 | 0.76 | NA | NA
| Analytical results for S in OREAS 461 (Certified Value 475 ppm) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 544 | 400 | 500 | 450 |  |  | 1
| 1 | 2 | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 469 | 300 | 550 | 450 |  |  | #N/A
| 1 | 3 | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 490 | 548 | 400 | 550 | 500 |  |  | 16
| 1 | 4 | 400 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 546 | 400 | 500 | 500 |  |  | 475.4111
| 1 | 5 | 400 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 551 | 400 | 500 | 500 |  |  | 105
| 1 | 6 | 500 | 500 | 500 | 500 | 500 | 500 | 400 | 480 | 359 | 400 | 550 | 500
| Mean |  | 466.6667 | 500 | 500 | 500 | 500 | 416.6667 | 400 | 481.6667 | 502.8333 | 383.3333 | 525 | 483.3333
| Median |  | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 545 | 400 | 525 | 500
| Std Dev. |  | 51.6398 | 0 | 0 | 0 | 0 | 40.8248 | 0 | 4.0825 | 77.1373 | 40.8248 | 27.3861 | 25.8199
| Rel.Std.Dev. |  | 11.07% | 0% | 0% | 0% | 0% | 9.798% | 0% | 0.8476% | 15.34% | 10.65% | 5.216% | 5.342%
| PDM3 |  | -1.839% | 5.172% | 5.172% | 5.172% | 5.172% | -12.36% | -15.86% | 1.316% | 5.768% | -19.37% | 10.43% | 1.666%
| Z-Score (Absolute) |  | 0.96 | 0.32 | 0.32 | 0.32 | 0.32 | 2.89 | 3.53 | 0.39 | 1.54 | 3.53 | 1.28 | 0.32
| Analytical results for Sb in OREAS 461 (Certified Value 2.42 ppm) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.35 | 2.67 | 2.58 | 2.49 | 2.43 | 2.1807 | 2.4 | 2.3 | 2.34 | 2.34 | 1.8 |  |  |  | 1
| 1 | 2 | 2.34 | 2.54 | 2.7 | 2.4 | 2.38 | 2.2331 | 2.5 | 2.5 | 2.33 | 2.12 | 1.8 |  |  |  | #N/A
| 1 | 3 | 2.25 | 2.64 | 2.75 | 2.48 | 2.43 | 2.2398 | 2.4 | 2.5 | 2.37 | 2.29 | 2 |  |  |  | 16
| 1 | 4 | 2.51 | 2.68 | 2.74 | 2.65 | 2.35 | 2.386 | 2.5 | 2.1 | 2.34 | 2.1 | 2 |  |  |  | 2.4174
| 1 | 5 | 2.33 | 2.55 | 2.72 | 2.62 | 2.41 | 2.2053 | 2.4 | 2.3 | 2.34 | 2.02 | 2 |  |  |  | 106
| 1 | 6 | 2.31 | 2.52 | 2.65 | 2.62 | 2.48 | 2.4325 | 2.5 | 2.5 | 2.41 | 2.09 | 1.8
| Mean |  | 2.3483 | 2.6 | 2.69 | 2.5433 | 2.4133 | 2.2796 | 2.45 | 2.3667 | 2.355 | 2.16 | 1.9
| Median |  | 2.335 | 2.595 | 2.71 | 2.555 | 2.42 | 2.2364 | 2.45 | 2.4 | 2.34 | 2.11 | 1.9
| Std Dev. |  | 0.086814 | 0.071274 | 0.064498 | 0.10053 | 0.045019 | 0.10367 | 0.054772 | 0.1633 | 0.030166 | 0.1257 | 0.10954
| Rel.Std.Dev. |  | 3.697% | 2.741% | 2.398% | 3.953% | 1.865% | 4.548% | 2.236% | 6.9% | 1.281% | 5.819% | 5.766%
| PDM3 |  | -2.857% | 7.554% | 11.28% | 5.21% | -0.1677% | -5.702% | 1.349% | -2.098% | -2.581% | -10.65% | -21.4%
| Z-Score (Absolute) |  | 0.39 | 1.81 | 2.5 | 1.37 | 0.36 | 0.67 | 0.65 | 0 | 0.09 | 1.6 | 3.61
| Analytical results for Sc in OREAS 461 (Certified Value 38.8 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 1
2
| 1 | 1 | 32.3 | 41.6 | 39.8 | 40.3 | 40.2 | 37.291 | 40 | 36 | 38.2 | 35.2283 | 45 | 38 |  |  | 1
| 1 | 2 | 38.2 | 40.5 | 40.5 | 39.2 | 38 | 37.2491 | 40 | 36 | 38.1 | 35.352 | 44 | 38 |  |  | 18
| 1 | 3 | 39 | 41.5 | 40.8 | 39.9 | 40.3 | 37.8361 | 40 | 37 | 37.5 | 35.1453 | 44 | 39 |  |  | 16
| 1 | 4 | 39.2 | 41 | 39.3 | 39.8 | 38 | 38.2016 | 40 | 37 | 37 | 35.2188 | 44 | 37 |  |  | 38.8492
| 1 | 5 | 35.8 | 41 | 41 | 41.1 | 39.2 | 38.427 | 40 | 37 | 36.4 | 35.3446 | 43 | 38 |  |  | 107
| 1 | 6 | 34 | 41 | 40.1 | 40.9 | 39.1 | 37.9736 | 39 | 37 | 37 | 35.0765 | 44 | 39
| Mean |  | 36.4167 | 41.1 | 40.25 | 40.2 | 39.1333 | 37.8297 | 39.8333 | 36.6667 | 37.3667 | 35.2276 | 44 | 38.1667
| Median |  | 37 | 41 | 40.3 | 40.1 | 39.15 | 37.9049 | 40 | 37 | 37.25 | 35.2235 | 44 | 38
| Std Dev. |  | 2.8555 | 0.4 | 0.64109 | 0.71554 | 1.0073 | 0.47814 | 0.40825 | 0.5164 | 0.70048 | 0.10847 | 0.63246 | 0.75277
| Rel.Std.Dev. |  | 7.841% | 0.9732% | 1.593% | 1.78% | 2.574% | 1.264% | 1.025% | 1.408% | 1.875% | 0.3079% | 1.437% | 1.972%
| PDM3 |  | -6.262% | 5.794% | 3.606% | 3.477% | 0.7313% | -2.624% | 2.533% | -5.618% | -3.816% | -9.322% | 13.26% | -1.757%
| Z-Score (Absolute) |  | 0.96 | 1.05 | 0.69 | 0.66 | 0.21 | 0.35 | 0.51 | 0.85 | 0.55 | 1.47 | 2.29 | 0.21
| Analytical results for Se in OREAS 461 (Indicative Value 6.48 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 4 | < 50 | 5 | 7 | 9 | 0.97792 | 9 | 6 | < 0.5 | < 0.5 | 5 |  |  |  | 1
| 1 | 2 | 4 | < 50 | 6 | 6 | 8 | 0.79493 | 10 | 8 | < 0.5 | < 0.5 | 5 |  |  |  | 15
| 1 | 3 | 4 | < 50 | 6 | 7 | 9 | 0.6892 | 10 | 6 | < 0.5 | < 0.5 | 5 |  |  |  | 16
| 1 | 4 | 4 | < 50 | 5 | 6 | 8 | 0.8 | 10 | 7 | < 0.5 | < 0.5 | 5 |  |  |  | 6.4762
| 1 | 5 | 4 | < 50 | 6 | 6 | 9 | 1.033 | 9 | 6 | < 0.5 | < 0.5 | 5 |  |  |  | 26
| 1 | 6 | 4 | < 50 | 5 | 6 | 8 | 0.86193 | 9 | 6 | < 0.5 | < 0.5 | 5
| Mean |  | 4 |  | 5.5 | 6.3333 | 8.5 | 0.8595 | 9.5 | 6.5 |  |  | 5
| Median |  | 4 |  | 5.5 | 6 | 8.5 | 0.83097 | 9.5 | 6 |  |  | 5
| Std Dev. |  | 0 |  | 0.54772 | 0.5164 | 0.54772 | 0.12716 | 0.54772 | 0.83666 |  |  | 0
| Rel.Std.Dev. |  | 0% |  | 9.959% | 8.154% | 6.444% | 14.79% | 5.766% | 12.87% |  |  | 0%
| PDM3 |  | -38.24% |  | -15.07% | -2.206% | 31.25% | -86.73% | 46.69% | 0.3676% |  |  | -22.79%
| Z-Score (Absolute) |  | 0.91 | NA | 0.2 | 0.2 | 1.23 | 2.41 | 1.71 | 0.28 | NA | NA | 0.44
| Analytical results for Sm in OREAS 461 (Certified Value 208 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 190.5 | 219 | 198.5 | 203 | 202 | 223.7505 | 207.5 | 200.8 | 215 | 206.45 | 213.9425 | 211.6 | 237 |  | 1
| 1 | 2 | 203 | 208 | 205 | 202 | 193.5 | 228.7411 | 204.2 | 212.3 | 216 | 203.43 | 213.105 | 209.8 | 223 |  | #N/A
| 1 | 3 | 202 | 225 | 208 | 206 | 200 | 221.665 | 209.7 | 206.2 | 214 | 209.18 | 213.4695 | 212.2 | 223 |  | 16
| 1 | 4 | 206 | 221 | 202 | 217 | 195.5 | 233.8646 | 210.7 | 205.5 | 209 | 215.51 | 212.3555 | 211.3 | 219 |  | 208.2366
| 1 | 5 | 196.5 | 210 | 208 | 219 | 197.5 | 233.7077 | 207.2 | 210.5 | 209 | 208.73 | 213.4181 | 212 | 217 |  | 108
| 1 | 6 | 188 | 225 | 206 | 212 | 202 | 228.2028 | 206.8 | 216.2 | 213 | 202.37 | 214.3532 | 215.8 | 235
| Mean |  | 197.6667 | 218 | 204.5833 | 209.8333 | 198.4167 | 228.322 | 207.6833 | 208.5833 | 212.6667 | 207.6117 | 213.4406 | 212.1167 | 225.6667
| Median |  | 199.25 | 220 | 205.5 | 209 | 198.75 | 228.472 | 207.35 | 208.35 | 213.5 | 207.59 | 213.4438 | 211.8 | 223
| Std Dev. |  | 7.2503 | 7.3756 | 3.7204 | 7.2503 | 3.5131 | 5.0026 | 2.2956 | 5.5003 | 3.0111 | 4.7414 | 0.68936 | 1.9944 | 8.3586
| Rel.Std.Dev. |  | 3.668% | 3.383% | 1.819% | 3.455% | 1.771% | 2.191% | 1.105% | 2.637% | 1.416% | 2.284% | 0.323% | 0.9402% | 3.704%
| PDM3 |  | -5.076% | 4.689% | -1.754% | 0.7668% | -4.716% | 9.645% | -0.2657% | 0.1665% | 2.127% | -0.3001% | 2.499% | 1.863% | 8.37%
| Z-Score (Absolute) |  | 2.27 | 1.53 | 0.98 | 0 | 2.13 | 3.46 | 0.4 | 0.23 | 0.53 | 0.42 | 0.67 | 0.43 | 2.96
| Analytical results for Sn in OREAS 461 (Certified Value 21.6 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 20.2 | 23.2 | 24.7 | 22.1 | 21.6 | 20.9812 | 22.8 | 19 | 17.7 | 23.5 | 21 |  |  |  | 1
| 1 | 2 | 20.2 | 22.3 | 25.6 | 21.5 | 21.2 | 19.707 | 22.2 | 19 | 17.3 | 23.4 | 21 |  |  |  | #N/A
| 1 | 3 | 19.8 | 23.4 | 25.9 | 22.4 | 22 | 19.6587 | 22.7 | 21 | 17.3 | 23.1 | 22 |  |  |  | 16
| 1 | 4 | 20.3 | 23.7 | 25.6 | 22.7 | 21.2 | 20.7433 | 22.1 | 19 | 17.9 | 24.1 | 23 |  |  |  | 21.6377
| 1 | 5 | 20.4 | 22.8 | 25.9 | 23.2 | 21.9 | 21.0801 | 21.5 | 20 | 17.8 | 23.2 | 22 |  |  |  | 109
| 1 | 6 | 20 | 22.9 | 25.3 | 22.8 | 22 | 20.2172 | 22.5 | 20 | 18 | 22.8 | 22
| Mean |  | 20.15 | 23.05 | 25.5 | 22.45 | 21.65 | 20.3979 | 22.3 | 19.6667 | 17.6667 | 23.35 | 21.8333
| Median |  | 20.2 | 23.05 | 25.6 | 22.55 | 21.75 | 20.4802 | 22.35 | 19.5 | 17.75 | 23.3 | 22
| Std Dev. |  | 0.21679 | 0.49295 | 0.45166 | 0.59582 | 0.37815 | 0.6295 | 0.47749 | 0.8165 | 0.30111 | 0.44159 | 0.75277
| Rel.Std.Dev. |  | 1.076% | 2.139% | 1.771% | 2.654% | 1.747% | 3.086% | 2.141% | 4.152% | 1.704% | 1.891% | 3.448%
| PDM3 |  | -6.875% | 6.527% | 17.85% | 3.754% | 0.0569% | -5.73% | 3.061% | -9.109% | -18.35% | 7.914% | 0.9042%
| Z-Score (Absolute) |  | 0.79 | 0.57 | 1.72 | 0.29 | 0.09 | 0.67 | 0.22 | 1.02 | 1.96 | 0.71 | 0
| Analytical results for Sr in OREAS 461 (Certified Value 564 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 601 | 600 | 518 | 576 | 535 | 573.4651 | 559 | 595.1 | 575 | 610.08 | 513.3074 | 369 | 585 | 516 | 1
| 1 | 2 | 619 | 595 | 519 | 581 | 547 | 575.7358 | 564.6 | 599.2 | 566 | 591.47 | 515.5238 | 364 | 585 | 504 | #N/A
| 1 | 3 | 592 | 602 | 517 | 583 | 520 | 578.9022 | 569.9 | 607 | 574 | 602.58 | 510.9785 | 375 | 585 | 513 | 16
| 1 | 4 | 607 | 603 | 518 | 576 | 501 | 565.6398 | 571.1 | 593.2 | 566 | 589.46 | 515.7509 | 371 | 580 | 508 | 564.1131
| 1 | 5 | 590 | 601 | 515 | 603 | 536 | 559.876 | 557.8 | 604.3 | 553 | 584.08 | 512.1413 | 360 | 590 | 508 | 110
| 1 | 6 | 591 | 605 | 518 | 596 | 505 | 568.3213 | 560.2 | 612 | 572 | 582.61 | 514.5032 | 376 | 580 | 513
| Mean |  | 600 | 601 | 517.5 | 585.8333 | 524 | 570.3234 | 563.7667 | 601.8 | 567.6667 | 593.38 | 513.7009 | 369.1667 | 584.1667 | 510.3333
| Median |  | 596.5 | 601.5 | 518 | 582 | 527.5 | 570.8932 | 562.4 | 601.75 | 569 | 590.465 | 513.9053 | 370 | 585 | 510.5
| Std Dev. |  | 11.4543 | 3.4059 | 1.3784 | 11.1609 | 18.4391 | 7.0406 | 5.712 | 7.2518 | 8.165 | 10.8169 | 1.9062 | 6.2423 | 3.7639 | 4.4121
| Rel.Std.Dev. |  | 1.909% | 0.5667% | 0.2664% | 1.905% | 3.519% | 1.235% | 1.013% | 1.205% | 1.438% | 1.823% | 0.3711% | 1.691% | 0.6443% | 0.8646%
| PDM3 |  | 6.362% | 6.539% | -8.263% | 3.85% | -7.111% | 1.101% | -0.06142% | 6.681% | 0.6299% | 5.188% | -8.937% | -34.56% | 3.555% | -9.534%
| Z-Score (Absolute) |  | 0.66 | 0.69 | 1.1 | 0.36 | 0.96 | 0.03 | 0.11 | 0.7 | 0.03 | 0.52 | 1.18 | 4.28 | 0.32 | 1.26
| Analytical results for Ta in OREAS 461 (Certified Value 20.9 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 22.9 | 23.6 | 21.9 | 21.8 | 18.6525 | 21.6 | 12.3 | 11.1 | 20.87 | 19.1 |  |  |  |  | 1
| 1 | 2 | 23.5 | 23 | 20.9 | 21.6 | 17.902 | 21.5 | 13.1 | 11.1 | 15.5 | 19.3 |  |  |  |  | #N/A
| 1 | 3 | 22.7 | 23.7 | 21.8 | 22.5 | 18.6311 | 22.6 | 10.8 | 10.8 | 17.62 | 21.2 |  |  |  |  | 16
| 1 | 4 | 22.7 | 22 | 22.1 | 21.5 | 18.8859 | 21 | 13 | 11.3 | 17.48 | 19.7 |  |  |  |  | 20.9304
| 1 | 5 | 22.2 | 23.5 | 22.5 | 22 | 20.3728 | 20.9 | 12.5 | 11.5 | 17.92 | 20.7 |  |  |  |  | 111
| 1 | 6 | 22.7 | 23 | 22.4 | 22.3 | 19.4891 | 20.7 | 12.5 | 11.5 | 17.75 | 20.1
| Mean |  | 22.7833 | 23.1333 | 21.9333 | 21.95 | 18.9889 | 21.3833 | 12.3667 | 11.2167 | 17.8567 | 20.0167
| Median |  | 22.7 | 23.25 | 22 | 21.9 | 18.7692 | 21.25 | 12.5 | 11.2 | 17.685 | 19.9
| Std Dev. |  | 0.42151 | 0.6314 | 0.57504 | 0.3937 | 0.84834 | 0.69113 | 0.82865 | 0.27142 | 1.7232 | 0.81588
| Rel.Std.Dev. |  | 1.85% | 2.729% | 2.622% | 1.794% | 4.468% | 3.232% | 6.701% | 2.42% | 9.65% | 4.076%
| PDM3 |  | 8.853% | 10.53% | 4.792% | 4.872% | -9.276% | 2.164% | -40.92% | -46.41% | -14.69% | -4.365%
| Z-Score (Absolute) |  | 0.74 | 0.86 | 0.44 | 0.44 | 0.61 | 0.24 | 2.85 | 3.37 | 1.22 | 0.24
| Analytical results for Tb in OREAS 461 (Certified Value 9.09 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 8.67 | 8.48 | 10.2 | 8.62 | 10 | 9.0451 | 9.1 | 8.75 | 11.2 | 9.59 | 13.8753 | 8.9 | 8.6 |  | 1
| 1 | 2 | 9.46 | 9.05 | 9.62 | 8.63 | 9.53 | 9.1344 | 9.1 | 9.08 | 11.1 | 9.3 | 13.9624 | 8.75 | 7.96 |  | #N/A
| 1 | 3 | 9.52 | 8.73 | 9.69 | 8.88 | 10.05 | 9.2887 | 9.3 | 8.99 | 11.1 | 9.4 | 14.2096 | 8.95 | 8.02 |  | 16
| 1 | 4 | 9.75 | 9.01 | 8.35 | 8.67 | 9.52 | 9.1004 | 9.2 | 8.7 | 11.1 | 9.7 | 14.1303 | 8.8 | 8.04 |  | 9.0942
| 1 | 5 | 9.41 | 8.62 | 9.77 | 8.7 | 10.1 | 9.2158 | 9.1 | 9.07 | 11.1 | 9.39 | 14.0595 | 8.75 | 8 |  | 112
| 1 | 6 | 9.09 | 8.78 | 9.32 | 9.05 | 9.93 | 9.0506 | 9.3 | 9.02 | 11.2 | 9.27 | 14.1682 | 9.05 | 8.48
| Mean |  | 9.3167 | 8.7783 | 9.4917 | 8.7583 | 9.855 | 9.1392 | 9.1833 | 8.935 | 11.1333 | 9.4417 | 14.0676 | 8.8667 | 8.1833
| Median |  | 9.435 | 8.755 | 9.655 | 8.685 | 9.965 | 9.1174 | 9.15 | 9.005 | 11.1 | 9.395 | 14.0949 | 8.85 | 8.03
| Std Dev. |  | 0.38156 | 0.22085 | 0.62736 | 0.17128 | 0.26174 | 0.096405 | 0.098319 | 0.1667 | 0.05164 | 0.16893 | 0.12831 | 0.12111 | 0.28012
| Rel.Std.Dev. |  | 4.095% | 2.516% | 6.61% | 1.956% | 2.656% | 1.055% | 1.071% | 1.866% | 0.4638% | 1.789% | 0.9121% | 1.366% | 3.423%
| PDM3 |  | 2.447% | -3.473% | 4.371% | -3.693% | 8.366% | 0.495% | 0.9805% | -1.75% | 22.42% | 3.821% | 54.69% | -2.502% | -10.02%
| Z-Score (Absolute) |  | 0.22 | 0.67 | 0.89 | 0.8 | 1.12 | 0.07 | 0 | 0.41 | 3.25 | 0.43 | 8.13 | 0.53 | 1.8
| Analytical results for Te in OREAS 461 (Certified Value 0.34 ppm) | [15 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.31 | 0.33 | 0.39 | 0.32 | 0.27 | 0.29853 | < 1 | < 0.5 | 0.37 | 0.3 | 0.2 |  |  |  | 1
| 1 | 2 | 0.26 | 0.37 | 0.44 | 0.32 | 0.31 | 0.40095 | < 1 | < 0.5 | 0.39 | 0.3 | 0.2 |  |  |  | 16
| 1 | 3 | 0.27 | 0.33 | 0.39 | 0.31 | 0.32 | 0.29853 | < 1 | < 0.5 | 0.42 | 0.3 | 0.2 |  |  |  | 16
| 1 | 4 | 0.27 | 0.29 | 0.49 | 0.31 | 0.3 | 0.3295 | < 1 | < 0.5 | 0.42 | 0.3 | 0.2 |  |  |  | 0.34127
| 1 | 5 | 0.31 | 0.36 | 0.44 | 0.35 | 0.31 | 0.40241 | < 1 | < 0.5 | 0.4 | 0.3 | 0.2 |  |  |  | 113
| 1 | 6 | 0.3 | 0.31 | 0.4 | 0.33 | 0.29 | 0.441 | < 1 | < 0.5 | 0.41 | 0.3 | 0.2
| Mean |  | 0.28667 | 0.33167 | 0.425 | 0.32333 | 0.3 | 0.36182 |  |  | 0.40167 | 0.3 | 0.2
| Median |  | 0.285 | 0.33 | 0.42 | 0.32 | 0.305 | 0.36522 |  |  | 0.405 | 0.3 | 0.2
| Std Dev. |  | 0.022509 | 0.029944 | 0.03937 | 0.015055 | 0.017889 | 0.060833 |  |  | 0.019408 | 0 | 3.0405e-17
| Rel.Std.Dev. |  | 7.852% | 9.028% | 9.264% | 4.656% | 5.963% | 16.81% |  |  | 4.832% | 0% | 1.52e-14%
| PDM3 |  | -16% | -2.814% | 24.54% | -5.256% | -12.09% | 6.022% |  |  | 17.7% | -12.09% | -41.4%
| Z-Score (Absolute) |  | 0.67 | 0.15 | 1.87 | 0 | 0.43 | 0.71 | NA | NA | 1.44 | 0.43 | 2.27
| Analytical results for Th in OREAS 461 (Certified Value 207 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 172.5 | 195 | 216 | 188.5 | 208 | 197.5202 | 199 | 215 | 196.86 | 214.71 | 198 |  |  |  | 1
| 1 | 2 | 210 | 212 | 223 | 189 | 199 | 197.4394 | 204.8 | 219 | 196.22 | 216.5911 | 201 |  |  |  | #N/A
| 1 | 3 | 216 | 214 | 225 | 193 | 212 | 200.1387 | 202.3 | 219 | 203.54 | 215.0865 | 203 |  |  |  | 16
| 1 | 4 | 219 | 211 | 218 | 198.5 | 202 | 204.8934 | 201.2 | 210 | 204.43 | 215.1762 | 203 |  |  |  | 206.55
| 1 | 5 | 187 | 210 | 225 | 197.5 | 204 | 204.727 | 205.9 | 215 | 199.4 | 214.8293 | 203 |  |  |  | 114
| 1 | 6 | 181.5 | 209 | 221 | 204 | 207 | 199.2399 | 207 | 216 | 198.33 | 215.4659 | 213
| Mean |  | 197.6667 | 208.5 | 221.3333 | 195.0833 | 205.3333 | 200.6598 | 203.3667 | 215.6667 | 199.7967 | 215.3098 | 203.5
| Median |  | 198.5 | 210.5 | 222 | 195.25 | 205.5 | 199.6893 | 203.55 | 215.5 | 198.865 | 215.1314 | 203
| Std Dev. |  | 19.7577 | 6.8337 | 3.7238 | 6.0284 | 4.6332 | 3.3758 | 3.052 | 3.3267 | 3.4411 | 0.68183 | 5.0498
| Rel.Std.Dev. |  | 9.995% | 3.278% | 1.682% | 3.09% | 2.256% | 1.682% | 1.501% | 1.543% | 1.722% | 0.3167% | 2.481%
| PDM3 |  | -4.301% | 0.9441% | 7.157% | -5.552% | -0.589% | -2.852% | -1.541% | 4.414% | -3.27% | 4.241% | -1.477%
| Z-Score (Absolute) |  | 0.13 | 1.48 | 3.39 | 1.56 | 0.37 | 0.51 | 0 | 2.32 | 0.67 | 2.26 | 0.33
| Analytical results for Ti in OREAS 461 (Certified Value 1.06 wt.%) | [15 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 1.12 | 1.155 | 1.1 | 1.015 | 1.005 | 1.032 | 1.03 | 0.6646 | 0.848 | 1.2764 | 0.66 | 1.33 |  |  | 1
| 1 | 2 | 1.17 | 1.125 | 1.085 | 0.973 | 1.07 | 0.95489 | 0.99 | 0.6268 | 0.881 | 1.2904 | 0.66 | 1.32 |  |  | #N/A
| 1 | 3 | 1.135 | 1.17 | 1.065 | 1.02 | 1.035 | 0.98844 | 1.01 | 0.6967 | 0.877 | 1.2475 | 0.65 | 1.45 |  |  | 16
| 1 | 4 | 1.105 | 1.24 | 1.08 | 1 | 0.991 | 0.93209 | 0.94 | 0.6788 | 0.856 | 1.2754 | 0.66 | 1.42 |  |  | 1.0566
| 1 | 5 | 1.095 | 1.19 | 1.085 | 1.04 | 1.065 | 0.95876 | 0.93 | 0.689 | 0.854 | 1.2401 | 0.67 | 1.4 |  |  | 115
| 1 | 6 | 1.135 | 1.175 | 1.105 | 1.06 | 1.015 | 1.022 | 0.93 | 0.6851 | 0.851 | 1.216 | 0.67 | 1.38
| Mean |  | 1.1267 | 1.1758 | 1.0867 | 1.018 | 1.0302 | 0.98137 | 0.97167 | 0.6735 | 0.86117 | 1.2576 | 0.66167 | 1.3833
| Median |  | 1.1275 | 1.1725 | 1.085 | 1.0175 | 1.025 | 0.9736 | 0.965 | 0.68195 | 0.855 | 1.2614 | 0.66 | 1.39
| Std Dev. |  | 0.026583 | 0.038395 | 0.014376 | 0.030364 | 0.032314 | 0.039773 | 0.044008 | 0.025303 | 0.014134 | 0.027874 | 0.0075277 | 0.050859
| Rel.Std.Dev. |  | 2.359% | 3.265% | 1.323% | 2.983% | 3.137% | 4.053% | 4.529% | 3.757% | 1.641% | 2.216% | 1.138% | 3.677%
| PDM3 |  | 6.634% | 11.29% | 2.848% | -3.651% | -2.499% | -7.118% | -8.036% | -36.26% | -18.49% | 19.03% | -37.38% | 30.93%
| Z-Score (Absolute) |  | 0.54 | 0.8 | 0.33 | 0.03 | 0.03 | 0.23 | 0.28 | 1.81 | 0.86 | 1.24 | 1.92 | 1.9
| Analytical results for Tl in OREAS 461 (Certified Value 0.1 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.11 | 0.16 | 0.11 | 0.12 | 0.11 | 0.097011 | 0.1 | < 0.2 | 0.1 | 0.1 | 0.22592 | 0.1 | < 0.1 |  | 1
| 1 | 2 | 0.12 | 0.21 | 0.1 | 0.11 | 0.11 | 0.095978 | 0.12 | < 0.2 | 0.11 | 0.1 | 0.20547 | 0.1 | 0.1 |  | 17
| 1 | 3 | 0.11 | 0.23 | 0.11 | 0.1 | 0.11 | 0.10073 | 0.1 | < 0.2 | 0.1 | 0.1 | 0.22987 | 0.1 | < 0.1 |  | 16
| 1 | 4 | 0.12 | 0.22 | 0.1 | 0.1 | 0.09 | 0.11122 | 0.11 | < 0.2 | 0.1 | 0.09 | 0.19515 | 0.1 | < 0.1 |  | 0.10384
| 1 | 5 | 0.11 | 0.21 | 0.11 | 0.11 | 0.11 | 0.11527 | 0.1 | < 0.2 | 0.1 | 0.09 | 0.2047 | 0.1 | < 0.1 |  | 116
| 1 | 6 | 0.11 | 0.19 | 0.11 | 0.11 | 0.1 | 0.11994 | 0.08 | < 0.2 | 0.1 | 0.09 | 0.18985 | 0.1 | 0.1
| Mean |  | 0.11333 | 0.20333 | 0.10667 | 0.10833 | 0.105 | 0.10669 | 0.10167 |  | 0.10167 | 0.095 | 0.20849 | 0.1 | 0.1
| Median |  | 0.11 | 0.21 | 0.11 | 0.11 | 0.11 | 0.10598 | 0.1 |  | 0.1 | 0.095 | 0.20509 | 0.1 | 0.1
| Std Dev. |  | 0.005164 | 0.025033 | 0.005164 | 0.0075277 | 0.0083666 | 0.010137 | 0.013292 |  | 0.0040825 | 0.0054772 | 0.016183 | 1.5202e-17 | 0
| Rel.Std.Dev. |  | 4.556% | 12.31% | 4.841% | 6.949% | 7.968% | 9.501% | 13.07% |  | 4.016% | 5.766% | 7.762% | 1.52e-14% | 0%
| PDM3 |  | 9.147% | 95.82% | 2.726% | 4.331% | 1.121% | 2.751% | -2.089% |  | -2.089% | -8.51% | 100.8% | -3.694% | -3.694%
| Z-Score (Absolute) |  | 1.01 | 13.15 | 0.11 | 0.34 | 0.11 | 0.12 | 0.56 | NA | 0.56 | 1.46 | 13.85 | 0.79 | 0.79
| Analytical results for Tm in OREAS 461 (Certified Value 0.7 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.75 | 0.72 | 0.74 | 0.68 | 0.72 | 0.7271 | 0.6543 | 0.8 | 0.66 | 0.61 | 0.85699 | 0.65 | 0.7 |  | 1
| 1 | 2 | 0.82 | 0.71 | 0.74 | 0.68 | 0.69 | 0.75963 | 0.69235 | 0.7 | 0.67 | 0.63 | 0.84328 | 0.65 | 0.64 |  | #N/A
| 1 | 3 | 0.83 | 0.73 | 0.74 | 0.7 | 0.72 | 0.71593 | 0.69481 | 0.7 | 0.67 | 0.63 | 0.85542 | 0.65 | 0.7 |  | 16
| 1 | 4 | 0.81 | 0.73 | 0.67 | 0.69 | 0.7 | 0.71171 | 0.69522 | 0.7 | 0.66 | 0.65 | 0.85063 | 0.65 | 0.7 |  | 0.70061
| 1 | 5 | 0.77 | 0.73 | 0.75 | 0.69 | 0.72 | 0.73368 | 0.64421 | 0.7 | 0.67 | 0.65 | 0.87224 | 0.65 | 0.68 |  | 117
| 1 | 6 | 0.76 | 0.72 | 0.71 | 0.7 | 0.74 | 0.75799 | 0.69714 | 0.7 | 0.67 | 0.61 | 0.89374 | 0.65 | 0.7
| Mean |  | 0.79 | 0.72333 | 0.725 | 0.69 | 0.715 | 0.73434 | 0.67967 | 0.71667 | 0.66667 | 0.63 | 0.86205 | 0.65 | 0.68667
| Median |  | 0.79 | 0.725 | 0.74 | 0.69 | 0.72 | 0.73039 | 0.69358 | 0.7 | 0.67 | 0.63 | 0.8562 | 0.65 | 0.7
| Std Dev. |  | 0.034059 | 0.008165 | 0.030166 | 0.0089443 | 0.017607 | 0.020507 | 0.023825 | 0.040825 | 0.005164 | 0.017889 | 0.018224 | 0 | 0.024221
| Rel.Std.Dev. |  | 4.311% | 1.129% | 4.161% | 1.296% | 2.462% | 2.793% | 3.505% | 5.696% | 0.7746% | 2.839% | 2.114% | 0% | 3.527%
| PDM3 |  | 12.76% | 3.243% | 3.481% | -1.515% | 2.054% | 4.814% | -2.988% | 2.292% | -4.845% | -10.08% | 23.04% | -7.224% | -1.99%
| Z-Score (Absolute) |  | 1.78 | 0.2 | 0.24 | 0.59 | 0 | 0.46 | 0.84 | 0.04 | 1.15 | 2.02 | 3.5 | 1.55 | 0.67
| Analytical results for U in OREAS 461 (Certified Value 4.4 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.4 | 4.2 | 4.5 | 4.2 | 4.5 | 4.2287 | 4.41 | 4.1 | 4.37 | 4.32 | 5.0784 | 4.1 | 3.7 |  | 1
| 1 | 2 | 4.4 | 4.4 | 4.6 | 4.2 | 4.4 | 4.2613 | 4.37 | 4.2 | 4.44 | 4.17 | 4.7527 | 4 | 3.8 |  | #N/A
| 1 | 3 | 4.4 | 4.4 | 4.8 | 4.3 | 4.6 | 4.4315 | 4.38 | 4.1 | 4.44 | 4.34 | 5.0481 | 4.1 | 3.8 |  | 16
| 1 | 4 | 4.5 | 4.5 | 4.6 | 4.5 | 4.4 | 4.4749 | 4.23 | 4 | 4.38 | 4.43 | 4.8021 | 4.1 | 3.8 |  | 4.4013
| 1 | 5 | 4.3 | 4.4 | 4.7 | 4.4 | 4.5 | 4.6947 | 4.41 | 4.2 | 4.55 | 4.62 | 4.973 | 4 | 3.8 |  | 118
| 1 | 6 | 4.5 | 4.3 | 4.7 | 4.5 | 4.5 | 4.5023 | 4.39 | 4.3 | 4.45 | 4.12 | 5.1162 | 4.2 | 4.1
| Mean |  | 4.4167 | 4.3667 | 4.65 | 4.35 | 4.4833 | 4.4322 | 4.365 | 4.15 | 4.4383 | 4.3333 | 4.9618 | 4.0833 | 3.8333
| Median |  | 4.4 | 4.4 | 4.65 | 4.35 | 4.5 | 4.4532 | 4.385 | 4.15 | 4.44 | 4.33 | 5.0106 | 4.1 | 3.8
| Std Dev. |  | 0.075277 | 0.10328 | 0.10488 | 0.13784 | 0.075277 | 0.17105 | 0.068044 | 0.10488 | 0.064317 | 0.18107 | 0.15116 | 0.075277 | 0.13663
| Rel.Std.Dev. |  | 1.704% | 2.365% | 2.256% | 3.169% | 1.679% | 3.859% | 1.559% | 2.527% | 1.449% | 4.179% | 3.046% | 1.844% | 3.564%
| PDM3 |  | 0.3501% | -0.7859% | 5.652% | -1.165% | 1.865% | 0.7038% | -0.8238% | -5.709% | 0.8424% | -1.543% | 12.74% | -7.223% | -12.9%
| Z-Score (Absolute) |  | 0.28 | 0.29 | 2.97 | 0.48 | 1.05 | 0.46 | 0 | 2.78 | 0.53 | 0.67 | 6.55 | 3.55 | 7.03
| Analytical results for V in OREAS 461 (Certified Value 352 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 341 | 372 | 297 | 326 | 310 | 353.9416 | 342.7 | 363 | 352 | 392 | 375.88 | 165 | 370 |  | 1
| 1 | 2 | 340 | 370 | 299 | 330 | 322 | 355.6957 | 354.2 | 361 | 356 | 402 | 374.834 | 164 | 370 |  | #N/A
| 1 | 3 | 330 | 376 | 294 | 329 | 309 | 353.0633 | 354 | 367 | 351 | 397 | 376.89 | 162 | 380 |  | 16
| 1 | 4 | 327 | 379 | 298 | 328 | 299 | 372.3048 | 366.4 | 364 | 351 | 410 | 374.988 | 164 | 360 |  | 352.0608
| 1 | 5 | 326 | 377 | 300 | 337 | 318 | 367.734 | 366.7 | 365 | 352 | 394 | 376.861 | 161 | 365 |  | 119
| 1 | 6 | 332 | 372 | 300 | 336 | 297 | 366.6284 | 366.9 | 370 | 341 | 394 | 375.654 | 165 | 375
| Mean |  | 332.6667 | 374.3333 | 298 | 331 | 309.1667 | 361.5613 | 358.4833 | 365 | 350.5 | 398.1667 | 375.8512 | 163.5 | 370
| Median |  | 331 | 374 | 298.5 | 329.5 | 309.5 | 361.1621 | 360.3 | 364.5 | 351.5 | 395.5 | 375.767 | 164 | 370
| Std Dev. |  | 6.4395 | 3.5024 | 2.2804 | 4.4721 | 9.9482 | 8.2931 | 9.8852 | 3.1623 | 5.01 | 6.7651 | 0.88525 | 1.6432 | 7.0711
| Rel.Std.Dev. |  | 1.936% | 0.9356% | 0.7652% | 1.351% | 3.218% | 2.294% | 2.758% | 0.8664% | 1.429% | 1.699% | 0.2355% | 1.005% | 1.911%
| PDM3 |  | -5.509% | 6.326% | -15.36% | -5.982% | -12.18% | 2.699% | 1.824% | 3.675% | -0.4433% | 13.1% | 6.757% | -53.56% | 5.095%
| Z-Score (Absolute) |  | 1 | 0.62 | 2.35 | 1.07 | 1.91 | 0.12 | 0 | 0.25 | 0.31 | 1.54 | 0.67 | 7.57 | 0.45
| Analytical results for W in OREAS 461 (Certified Value 2.5 ppm) | [15 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.6 | 2.7 | 3.3 | 2.7 | 2.8 | 2.4648 | 2 | 2 | 2.2 | 2.4 | 2.5 |  |  |  | 1
| 1 | 2 | 2.9 | 4 | 3.2 | 2.7 | 2.7 | 2.2734 | 2 | 2 | 2.1 | 2.2 | 2.5 |  |  |  | #N/A
| 1 | 3 | 2.7 | 2.7 | 3.3 | 2.7 | 2.8 | 2.3928 | 2 | 2 | 2.2 | 2.4 | 2.5 |  |  |  | 16
| 1 | 4 | 2.8 | 2.8 | 3 | 2.7 | 2.7 | 2.4267 | 2 | 2 | 2.4 | 2.2 | 2.5 |  |  |  | 2.4971
| 1 | 5 | 2.6 | 2.7 | 3.2 | 2.8 | 2.8 | 2.3948 | 2 | 2 | 2.2 | 2.2 | 2.5 |  |  |  | 120
| 1 | 6 | 2.7 | 2.7 | 3.1 | 2.8 | 2.8 | 2.3342 | 2 | 2 | 2.3 | 2.2 | 2.5
| Mean |  | 2.7167 | 2.9333 | 3.1833 | 2.7333 | 2.7667 | 2.3811 | 2 | 2 | 2.2333 | 2.2667 | 2.5
| Median |  | 2.7 | 2.7 | 3.2 | 2.7 | 2.8 | 2.3938 | 2 | 2 | 2.2 | 2.2 | 2.5
| Std Dev. |  | 0.1169 | 0.52409 | 0.1169 | 0.05164 | 0.05164 | 0.068145 | 0 | 0 | 0.10328 | 0.10328 | 0
| Rel.Std.Dev. |  | 4.303% | 17.87% | 3.672% | 1.889% | 1.866% | 2.862% | 0% | 0% | 4.624% | 4.556% | 0%
| PDM3 |  | 8.794% | 17.47% | 27.48% | 9.462% | 10.8% | -4.644% | -19.91% | -19.91% | -10.56% | -9.227% | 0.1173%
| Z-Score (Absolute) |  | 0.63 | 0.64 | 1.97 | 0.67 | 0.77 | 0.34 | 1.44 | 1.44 | 0.87 | 0.67 | 0
| Analytical results for Y in OREAS 461 (Certified Value 83 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 76.3 | 83 | 89.3 | 86.4 | 86.7 | 78.6506 | 83.9 | 80.8 | 81.5 | 87.06 | 84.3177 | 77.1 | 79.7 |  | 1
| 1 | 2 | 84.8 | 77.1 | 87.5 | 85 | 82.1 | 79.3019 | 84 | 82.1 | 79.5 | 87.87 | 84.2674 | 76.7 | 84.4 |  | #N/A
| 1 | 3 | 85.9 | 80.8 | 90.1 | 86.9 | 86 | 83.9277 | 85.4 | 81.8 | 81.9 | 87.79 | 83.6995 | 77.4 | 81.4 |  | 16
| 1 | 4 | 89.6 | 80.2 | 85.2 | 88.2 | 82.7 | 83.3788 | 88.6 | 80 | 79.9 | 85.96 | 84.2908 | 77.4 | 80.7 |  | 83.097
| 1 | 5 | 80.6 | 77.6 | 90.5 | 89.1 | 85.9 | 80.175 | 84.8 | 82 | 78.8 | 85.72 | 85.1148 | 76.9 | 82 |  | 121
| 1 | 6 | 78.9 | 78.1 | 84.4 | 88.3 | 86.3 | 78.9576 | 86 | 83.5 | 79.5 | 84.62 | 85.9227 | 77.7 | 101
| Mean |  | 82.6833 | 79.4667 | 87.8333 | 87.3167 | 84.95 | 80.7319 | 85.45 | 81.7 | 80.1833 | 86.5033 | 84.6022 | 77.2 | 84.8667
| Median |  | 82.7 | 79.15 | 88.4 | 87.55 | 85.95 | 79.7384 | 85.1 | 81.9 | 79.7 | 86.51 | 84.3042 | 77.25 | 81.7
| Std Dev. |  | 4.941 | 2.2695 | 2.5781 | 1.5039 | 2.0037 | 2.3261 | 1.741 | 1.2 | 1.2336 | 1.2876 | 0.78882 | 0.36878 | 8.0597
| Rel.Std.Dev. |  | 5.976% | 2.856% | 2.935% | 1.722% | 2.359% | 2.881% | 2.037% | 1.469% | 1.538% | 1.488% | 0.9324% | 0.4777% | 9.497%
| PDM3 |  | -0.4978% | -4.369% | 5.7% | 5.078% | 2.23% | -2.846% | 2.832% | -1.681% | -3.506% | 4.099% | 1.811% | -7.097% | 2.13%
| Z-Score (Absolute) |  | 0 | 0.87 | 1.39 | 1.25 | 0.61 | 0.53 | 0.75 | 0.27 | 0.67 | 1.03 | 0.52 | 1.48 | 0.28
| Analytical results for Yb in OREAS 461 (Certified Value 3.06 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.09 | 2.95 | 3.06 | 2.88 | 3.02 | 2.9613 | 2.9006 | 2.9 | 3.81 | 3.02 | 3.4615 | 3 | 3.2 |  | 1
| 1 | 2 | 3.36 | 2.97 | 3.12 | 2.89 | 2.94 | 2.9617 | 2.8247 | 2.9 | 3.84 | 3.17 | 3.2961 | 3 | 3.35 |  | #N/A
| 1 | 3 | 3.34 | 3.01 | 3.05 | 2.95 | 3.06 | 3.3061 | 2.9131 | 3 | 3.81 | 3.04 | 3.5337 | 3.05 | 3.3 |  | 16
| 1 | 4 | 3.36 | 3.01 | 3.04 | 2.93 | 2.95 | 3.1299 | 3.0184 | 3 | 3.75 | 3.18 | 3.4511 | 2.95 | 3.25 |  | 3.0631
| 1 | 5 | 3.18 | 3.01 | 3.16 | 2.92 | 3.02 | 3.33 | 2.9369 | 2.9 | 3.79 | 3.15 | 3.6928 | 3.05 | 3.3 |  | 122
| 1 | 6 | 3.07 | 2.98 | 3.06 | 3 | 3.08 | 3.3169 | 2.8287 | 3.1 | 3.92 | 3.95 | 2.8579 | 3 | 3.45
| Mean |  | 3.2333 | 2.9883 | 3.0817 | 2.9283 | 3.0117 | 3.1677 | 2.9037 | 2.9667 | 3.82 | 3.2517 | 3.3822 | 3.0083 | 3.3083
| Median |  | 3.26 | 2.995 | 3.06 | 2.925 | 3.02 | 3.218 | 2.9069 | 2.95 | 3.81 | 3.16 | 3.4563 | 3 | 3.3
| Std Dev. |  | 0.13677 | 0.025626 | 0.047504 | 0.043551 | 0.056716 | 0.1756 | 0.072409 | 0.08165 | 0.057271 | 0.34879 | 0.28734 | 0.037639 | 0.08612
| Rel.Std.Dev. |  | 4.23% | 0.8575% | 1.542% | 1.487% | 1.883% | 5.544% | 2.494% | 2.752% | 1.499% | 10.73% | 8.496% | 1.251% | 2.603%
| PDM3 |  | 5.557% | -2.442% | 0.6052% | -4.401% | -1.68% | 3.413% | -5.203% | -3.149% | 24.71% | 6.155% | 10.42% | -1.789% | 8.005%
| Z-Score (Absolute) |  | 1.11 | 0.52 | 0 | 0.91 | 0.36 | 0.67 | 1.08 | 0.66 | 5 | 0.31 | 2.79 | 0.38 | 1.61
| Analytical results for Zn in OREAS 461 (Certified Value 139 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 138 | 142 | 119 | 120 | 119 | 147.3384 | 145.4541 | 142 | 156 | 148 | 222.899 | 85 | 145 | 142 | 1
| 1 | 2 | 137 | 141 | 119 | 120 | 124 | 151.7959 | 145.1305 | 153 | 157 | 152 | 223.209 | 87 | 145 | 132 | #N/A
| 1 | 3 | 133 | 143 | 117 | 122 | 120 | 148.195 | 145.7471 | 171 | 157 | 145 | 223.033 | 87 | 145 | 144 | 16
| 1 | 4 | 133 | 147 | 118 | 122 | 115 | 149.7978 | 144.214 | 145 | 157 | 147 | 221.409 | 89 | 145 | 132 | 139.3186
| 1 | 5 | 131 | 143 | 119 | 126 | 124 | 152.1224 | 143.1699 | 150 | 158 | 145 | 221.007 | 87 | 140 | 138 | 123
| 1 | 6 | 133 | 142 | 119 | 125 | 115 | 148.1167 | 144.6549 | 163 | 158 | 143 | 222.808 | 88 | 145 | 144
| Mean |  | 134.1667 | 143 | 118.5 | 122.5 | 119.5 | 149.561 | 144.7284 | 154 | 157.1667 | 146.6667 | 222.3942 | 87.1667 | 144.1667 | 138.6667
| Median |  | 133 | 142.5 | 119 | 122 | 119.5 | 148.9964 | 144.8927 | 151.5 | 157 | 146 | 222.8535 | 87 | 145 | 140
| Std Dev. |  | 2.7142 | 2.0976 | 0.83666 | 2.51 | 4.0373 | 2.025 | 0.94074 | 11.0635 | 0.75277 | 3.1411 | 0.93733 | 1.3292 | 2.0412 | 5.6095
| Rel.Std.Dev. |  | 2.023% | 1.467% | 0.706% | 2.049% | 3.379% | 1.354% | 0.65% | 7.184% | 0.479% | 2.142% | 0.4215% | 1.525% | 1.416% | 4.045%
| PDM3 |  | -3.698% | 2.642% | -14.94% | -12.07% | -14.23% | 7.352% | 3.883% | 10.54% | 12.81% | 5.274% | 59.63% | -37.43% | 3.48% | -0.4679%
| Z-Score (Absolute) |  | 0.61 | 0.07 | 1.68 | 1.41 | 1.61 | 0.43 | 0.11 | 0.74 | 0.95 | 0.24 | 5.39 | 3.81 | 0.07 | 0.31
| Analytical results for Zr in OREAS 461 (Certified Value 270 ppm) | [15 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 291 | 287 | 272 | 264 | 284 | 244.2632 | 219 | 325 | 264.2 | 194 | 245 | 284 |  |  | 1
| 1 | 2 | 309 | 289 | 265 | 256 | 279 | 242.4219 | 210 | 321 | 271.6 | 194 | 240 | 289 |  |  | #N/A
| 1 | 3 | 300 | 276 | 301 | 264 | 276 | 247.4138 | 216 | 339 | 263.6 | 209 | 240 | 282 |  |  | 16
| 1 | 4 | 292 | 292 | 261 | 263 | 267 | 265.4465 | 214 | 324 | 258.2 | 200 | 230 | 280 |  |  | 269.8554
| 1 | 5 | 287 | 274 | 308 | 268 | 277 | 265.4442 | 220 | 327 | 253.9 | 180 | 235 | 276 |  |  | 124
| 1 | 6 | 290 | 278 | 265 | 268 | 275 | 254.7995 | 217 | 341 | 258.9 | 194 | 240 | 280
| Mean |  | 294.8333 | 282.6667 | 278.6667 | 263.8333 | 276.3333 | 253.2982 | 216 | 329.5 | 261.7333 | 195.1667 | 238.3333 | 281.8333
| Median |  | 291.5 | 282.5 | 268.5 | 264 | 276.5 | 251.1067 | 216.5 | 326 | 261.25 | 194 | 240 | 281
| Std Dev. |  | 8.1833 | 7.5807 | 20.4418 | 4.4008 | 5.5737 | 10.3142 | 3.6332 | 8.3845 | 6.1448 | 9.4745 | 5.164 | 4.4008
| Rel.Std.Dev. |  | 2.776% | 2.682% | 7.336% | 1.668% | 2.017% | 4.072% | 1.682% | 2.545% | 2.348% | 4.855% | 2.167% | 1.561%
| PDM3 |  | 9.256% | 4.747% | 3.265% | -2.232% | 2.401% | -6.136% | -19.96% | 22.1% | -3.01% | -27.68% | -11.68% | 4.439%
| Z-Score (Absolute) |  | 1.01 | 0.58 | 0.39 | 0.29 | 0.29 | 0.77 | 2.48 | 2.73 | 0.38 | 3.44 | 1.46 | 0.54
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.