## Abbreviations
Table 1. Abbreviations used for OREAS 461
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 461
Lab Slot | Laboratory Name and Location
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
Nagrom, Perth, WA, Australia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
SGS Vostok Limited, Chita, Russian Federation
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 461
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate Fusion XRF
Ce, Cerium (ppm) | 3562.359 | 44.7088 | 3525.6197 | 3599.0983 | IND | IND
Fe, Iron (wt.%) | 32.2378 | 0.54074 | 31.5187 | 32.9569 | 32.024 | 32.4515
La, Lanthanum (ppm) | 2720.4541 | 54.6493 | 2684.841 | 2756.0672 | 2603.578 | 2837.3302
Nd, Neodymium (ppm) | 1653.1237 | 113.7825 | 1544.0874 | 1762.16 | IND | IND
Pr, Praseodymium (ppm) | 483.9177 | 85.5539 | 400.3989 | 567.4366 | IND | IND
Th, Thorium (ppm) | 232.665 | 45.6019 | 170.8719 | 294.4581 | IND | IND
Thermogravimetry
LOI, Loss On Ignition @1000°C (wt.%) | 0.99639 | 0.175 | 0.80055 | 1.1922 | 0.9397 | 1.0531
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.9859 | 0.20778 | 5.839 | 6.1328 | 5.8848 | 6.087
Ba, Barium (ppm) | 928.8402 | 55.2115 | 898.0997 | 959.5808 | 912.0559 | 945.6245
Be, Beryllium (ppm) | 2.3019 | 0.376 | 1.911 | 2.6927 | IND | IND
Bi, Bismuth (ppm) | 2.4852 | 0.3867 | 2.0889 | 2.8815 | IND | IND
Ca, Calcium (wt.%) | 1.2585 | 0.057298 | 1.2216 | 1.2954 | 1.2302 | 1.2868
Ce, Cerium (ppm) | 3509.7977 | 131.9597 | 3436.6114 | 3582.9841 | 3413.6439 | 3605.9515
Co, Cobalt (ppm) | 11.5257 | 1.7734 | 9.6504 | 13.401 | 11.122 | 11.9294
Cr, Chromium (ppm) | 599.0322 | 40.4701 | 574.9298 | 623.1347 | 584.4418 | 613.6227
Cs, Cesium (ppm) | 0.78598 | 0.058996 | 0.74972 | 0.82224 | 0.74797 | 0.82399
Cu, Copper (ppm) | 60.1422 | 19.7586 | 40.6166 | 79.6679 | IND | IND
Dy, Dysprosium (ppm) | 34.8324 | 1.444 | 34.0291 | 35.6356 | 33.6646 | 36.0001
Er, Erbium (ppm) | 8.8026 | 0.29606 | 8.641 | 8.9641 | 8.4728 | 9.1323
Eu, Europium (ppm) | 46.7145 | 2.1891 | 45.4394 | 47.9895 | 45.0623 | 48.3666
Fe, Iron (wt.%) | 31.7733 | 1.3012 | 30.8208 | 32.7259 | 31.237 | 32.3097
Ga, Gallium (ppm) | 50.0389 | 5.8276 | 43.2986 | 56.7792 | 47.6607 | 52.4171
Gd, Gadolinium (ppm) | 100.3873 | 8.068 | 95.4177 | 105.3568 | 97.5016 | 103.2729
Hf, Hafnium (ppm) | 14.093 | 0.60698 | 13.7038 | 14.4821 | 13.4439 | 14.742
Ho, Holmium (ppm) | 4.5595 | 0.30588 | 4.3782 | 4.7408 | 4.3582 | 4.7608
In, Indium (ppm) | 0.60113 | 0.07596 | 0.52363 | 0.67862 | IND | IND
K, Potassium (wt.%) | 0.23655 | 0.033137 | 0.21254 | 0.26056 | 0.22427 | 0.24883
La, Lanthanum (ppm) | 2690.0867 | 157.5096 | 2601.3776 | 2778.7958 | 2634.7515 | 2745.422
Li, Lithium (ppm) | 12.0055 | 1.5327 | 10.3279 | 13.6832 | IND | IND
Lu, Lutetium (ppm) | 0.5154 | 0.043771 | 0.4895 | 0.5413 | 0.47708 | 0.55371
Mg, Magnesium (wt.%) | 1.047 | 0.051391 | 1.0128 | 1.0812 | 1.0231 | 1.0709
Mn, Manganese (ppm) | 652.6366 | 48.8555 | 615.979 | 689.2942 | 632.6331 | 672.6401
Mo, Molybdenum (ppm) | 45.6795 | 4.8335 | 41.0819 | 50.2771 | 43.3931 | 47.9659
Nb, Niobium (ppm) | 1296.3889 | 97.646 | 1231.4391 | 1361.3388 | 1260.1671 | 1332.6108
Nd, Neodymium (ppm) | 1628.7803 | 54.2871 | 1597.1832 | 1660.3774 | 1583.9924 | 1673.5682
Ni, Nickel (ppm) | 63.6945 | 5.4703 | 60.9591 | 66.4299 | 53.204 | 74.185
P, Phosphorus (wt.%) | 0.35872 | 0.019045 | 0.34433 | 0.37311 | 0.34358 | 0.37386
Pb, Lead (ppm) | 106.2928 | 6.3097 | 99.8829 | 112.7027 | 99.2991 | 113.2864
Pr, Praseodymium (ppm) | 488.5896 | 20.9195 | 477.1794 | 499.9998 | 473.1819 | 503.9974
Rb, Rubidium (ppm) | 13.5154 | 0.5424 | 13.2241 | 13.8068 | 13.0306 | 14.0003
S, Sulphur (ppm) | 517.6667 | 82.1688 | 432.951 | 602.3824 | IND | IND
Sb, Antimony (ppm) | 2.7613 | 0.21987 | 2.5311 | 2.9915 | 2.5308 | 2.9919
Si, Silicon (wt.%) | 14.7082 | 0.3598 | 14.3979 | 15.0184 | 14.4785 | 14.9378
Sm, Samarium (ppm) | 219.6468 | 7.3567 | 216.1211 | 223.1725 | 212.8418 | 226.4518
Sn, Tin (ppm) | 25.6303 | 3.3461 | 23.3942 | 27.8665 | 23.7379 | 27.5227
Sr, Strontium (ppm) | 578.9394 | 15.3214 | 570.071 | 587.8078 | 567.0484 | 590.8304
Ta, Tantalum (ppm) | 25.1318 | 1.0401 | 24.5824 | 25.6813 | 24.038 | 26.2256
Tb, Terbium (ppm) | 9.0752 | 0.30531 | 8.9201 | 9.2304 | 8.7589 | 9.3915
Th, Thorium (ppm) | 209.8552 | 6.8117 | 205.6201 | 214.0902 | 205.8514 | 213.859
Ti, Titanium (wt.%) | 1.8399 | 0.055067 | 1.8028 | 1.877 | 1.8051 | 1.8747
Tm, Thulium (ppm) | 0.88945 | 0.085451 | 0.84213 | 0.93677 | 0.83681 | 0.94209
U, Uranium (ppm) | 4.7882 | 0.19497 | 4.6866 | 4.8899 | 4.6361 | 4.9404
V, Vanadium (ppm) | 384.6194 | 16.4758 | 375.8163 | 393.4224 | 375.1275 | 394.1112
W, Tungsten (ppm) | 3.8778 | 0.66844 | 3.3536 | 4.4019 | IND | IND
Y, Yttrium (ppm) | 90.9849 | 4.3596 | 88.381 | 93.5887 | 88.4168 | 93.5529
Yb, Ytterbium (ppm) | 4.3941 | 0.24773 | 4.2659 | 4.5222 | 4.1409 | 4.6472
Zn, Zinc (ppm) | 161.7333 | 23.352 | 130.614 | 192.8526 | IND | IND
Zr, Zirconium (ppm) | 603.0104 | 26.8975 | 585.0611 | 620.9598 | 581.6316 | 624.3892
4-Acid Digestion
Ag, Silver (ppm) | < 2 | IND | IND | IND | IND | IND
Al, Aluminium (wt.%) | 5.9585 | 0.18013 | 5.8591 | 6.058 | 5.8563 | 6.0607
As, Arsenic (ppm) | 33.5537 | 1.0854 | 32.9617 | 34.1457 | 32.327 | 34.7804
Ba, Barium (ppm) | 957.0984 | 44.069 | 931.1315 | 983.0654 | 935.9555 | 978.2414
Be, Beryllium (ppm) | 2.0913 | 0.24371 | 1.9592 | 2.2234 | 1.9815 | 2.2011
Bi, Bismuth (ppm) | 2.4269 | 0.15827 | 2.3321 | 2.5217 | 2.3274 | 2.5264
Ca, Calcium (wt.%) | 1.2147 | 0.090313 | 1.162 | 1.2674 | 1.1958 | 1.2336
Cd, Cadmium (ppm) | < 0.1 | IND | IND | IND | IND | IND
Ce, Cerium (ppm) | 3508.4834 | 111.9702 | 3420.3035 | 3596.6632 | 3411.4065 | 3605.5602
Co, Cobalt (ppm) | 12.1248 | 0.44431 | 11.8106 | 12.439 | 11.7568 | 12.4928
Cr, Chromium (ppm) | 523.6841 | 73.1099 | 472.7326 | 574.6355 | 507.4963 | 539.8719
Cs, Cesium (ppm) | 0.76887 | 0.066983 | 0.72791 | 0.80982 | 0.73276 | 0.80497
Cu, Copper (ppm) | 55.8138 | 2.2885 | 54.6346 | 56.993 | 54.2832 | 57.3445
Dy, Dysprosium (ppm) | 33.5055 | 0.86045 | 33.117 | 33.8941 | 32.6607 | 34.3504
Er, Erbium (ppm) | 7.4597 | 0.25253 | 7.3063 | 7.613 | 7.2377 | 7.6816
Eu, Europium (ppm) | 47.6394 | 3.2615 | 45.7237 | 49.5551 | 46.5661 | 48.7127
Fe, Iron (wt.%) | 31.5178 | 0.95936 | 30.9495 | 32.0861 | 31.0755 | 31.9601
Ga, Gallium (ppm) | 45.5202 | 8.0263 | 37.719 | 53.3215 | 44.1388 | 46.9017
Gd, Gadolinium (ppm) | 97.5977 | 3.4838 | 95.4894 | 99.706 | 95.3525 | 99.8429
Hf, Hafnium (ppm) | 7.3912 | 0.83938 | 6.7803 | 8.002 | 6.9582 | 7.8241
Ho, Holmium (ppm) | 4.0576 | 0.11474 | 3.9904 | 4.1248 | 3.9529 | 4.1623
In, Indium (ppm) | 0.56812 | 0.027058 | 0.55328 | 0.58296 | 0.5396 | 0.59664
K, Potassium (wt.%) | 0.23312 | 0.0059742 | 0.23076 | 0.23549 | 0.22291 | 0.24334
La, Lanthanum (ppm) | 2476.481 | 227.9521 | 2337.8891 | 2615.0729 | 2422.1987 | 2530.7633
Li, Lithium (ppm) | 12.0397 | 1.3787 | 11.3262 | 12.7533 | 11.6348 | 12.4447
Lu, Lutetium (ppm) | 0.38585 | 0.065651 | 0.34607 | 0.42562 | 0.36197 | 0.40972
Mg, Magnesium (wt.%) | 0.99901 | 0.10801 | 0.9352 | 1.0628 | 0.97961 | 1.0184
Mn, Manganese (ppm) | 547.2361 | 79.7765 | 495.7591 | 598.7132 | 533.448 | 561.0243
Mo, Molybdenum (ppm) | 48.2464 | 2.5361 | 46.5606 | 49.9323 | 47.2489 | 49.244
Na, Sodium (wt.%) | 0.17817 | 0.01301 | 0.17029 | 0.18606 | IND | IND
Nd, Neodymium (ppm) | 1623.483 | 79.5193 | 1579.0041 | 1667.9619 | 1566.4025 | 1680.5636
Ni, Nickel (ppm) | 71.1258 | 3.1439 | 68.9999 | 73.2516 | 68.7203 | 73.5313
P, Phosphorus (wt.%) | 0.33785 | 0.02169 | 0.324 | 0.35171 | 0.33275 | 0.34296
Pb, Lead (ppm) | 103.9184 | 5.4831 | 99.811 | 108.0258 | 101.371 | 106.4658
Pr, Praseodymium (ppm) | 482.2815 | 20.1118 | 471.2553 | 493.3077 | 469.9093 | 494.6538
Rb, Rubidium (ppm) | 13.2859 | 0.98994 | 12.7563 | 13.8155 | 12.9543 | 13.6175
Re, Rhenium (ppm) | < 0.01 | IND | IND | IND | IND | IND
S, Sulphur (ppm) | 475.4111 | 48.6238 | 446.3341 | 504.4881 | IND | IND
Sb, Antimony (ppm) | 2.4174 | 0.1736 | 2.3039 | 2.5309 | 2.2985 | 2.5362
Sc, Scandium (ppm) | 38.8492 | 2.505 | 37.3175 | 40.381 | 37.7312 | 39.9673
Sm, Samarium (ppm) | 208.2366 | 7.4192 | 204.0846 | 212.3885 | 202.7799 | 213.6932
Sn, Tin (ppm) | 21.6377 | 2.0829 | 20.2223 | 23.053 | 20.7426 | 22.5328
Sr, Strontium (ppm) | 564.1131 | 35.1659 | 542.7011 | 585.5252 | 555.5121 | 572.7142
Ta, Tantalum (ppm) | 20.9304 | 1.96 | 19.2388 | 22.6219 | 19.707 | 22.1538
Tb, Terbium (ppm) | 9.0942 | 0.50289 | 8.7602 | 9.4281 | 8.8702 | 9.3182
Te, Tellurium (ppm) | 0.34127 | 0.055579 | 0.29901 | 0.38352 | 0.28728 | 0.39525
Th, Thorium (ppm) | 206.55 | 9.605 | 201.0808 | 212.0191 | 200.0615 | 213.0384
Ti, Titanium (wt.%) | 1.0566 | 0.11708 | 0.96494 | 1.1482 | 1.0112 | 1.1019
Tl, Thallium (ppm) | 0.10384 | 0.0084602 | 0.1001 | 0.10757 | IND | IND
Tm, Thulium (ppm) | 0.70061 | 0.046188 | 0.67342 | 0.7278 | 0.67679 | 0.72443
U, Uranium (ppm) | 4.4013 | 0.16104 | 4.3111 | 4.4914 | 4.2647 | 4.5378
V, Vanadium (ppm) | 352.0608 | 28.763 | 333.494 | 370.6275 | 345.0365 | 359.085
W, Tungsten (ppm) | 2.4971 | 0.36095 | 2.2506 | 2.7436 | 2.3476 | 2.6466
Y, Yttrium (ppm) | 83.097 | 3.7487 | 81.1129 | 85.0811 | 80.5917 | 85.6022
Yb, Ytterbium (ppm) | 3.0631 | 0.14929 | 2.9764 | 3.1498 | 2.9612 | 3.165
Zn, Zinc (ppm) | 139.3186 | 13.2856 | 130.99 | 147.6472 | 136.2141 | 142.423
Zr, Zirconium (ppm) | 269.8554 | 18.0276 | 256.8571 | 282.8536 | 256.8606 | 282.8501

## Indicative Values
Table 3. Indicative Values for OREAS 461
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Al | wt.% | 6.0814 | Lu | ppm | < 20 | Ta | ppm | < 100
Ba | ppm | 1000.1493 | Mg | wt.% | 1.0611 | Tb | ppm | < 20
Ca | wt.% | 1.2975 | Mn | ppm | 700.253 | Ti | wt.% | 1.8439
Cr | ppm | 429.9236 | Na | wt.% | 0.18558 | Tm | ppm | < 10
Dy | ppm | 40.661 | Nb | ppm | 1302.4525 | U | ppm | < 100
Er | ppm | 17.4902 | P | wt.% | 0.37099 | V | ppm | 407.9842
Eu | ppm | < 100 | Si | wt.% | 14.7966 | W | ppm | < 100
Gd | ppm | 86.7604 | Sm | ppm | 190.1345 | Y | ppm | 115.1665
Hf | ppm | < 100 | Sn | ppm | < 100 | Yb | ppm | 8.7819
Ho | ppm | 26.1894 | S | wt.% | 0.046786 | Zr | ppm | 592.2416
K | wt.% | 0.23576 | Sr | ppm | 422.7972
Thermogravimetry
H2O- | wt.% | 0.69667
Borate / Peroxide Fusion ICP
Ag | ppm | 15.1667 | Ge | ppm | 2.5 | Se | ppm | < 20
As | ppm | 79.3256 | Na | wt.% | 0.16624 | Te | ppm | < 1
B | ppm | 150.8333 | Re | ppm | < 0.1 | Tl | ppm | < 0.5
Cd | ppm | < 1 | Sc | ppm | 44
4-Acid Digestion
Ge | ppm | 3.5955 | Nb | ppm | 1152.7857 | Se | ppm | 6.4762

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 461
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Ce, ppm | 3562.359 | 44.7088 | 3472.9414 | 3651.7766 | 3428.2326 | 3696.4854 | 1.255% | 2.51% | 3.765% | 3384.2411 | 3740.477
Fe, wt.% | 32.2378 | 0.54074 | 31.1563 | 33.3192 | 30.6155 | 33.86 | 1.677% | 3.355% | 5.032% | 30.6259 | 33.8497
La, ppm | 2720.4541 | 54.6493 | 2611.1554 | 2829.7528 | 2556.5061 | 2884.4021 | 2.009% | 4.018% | 6.026% | 2584.4314 | 2856.4768
Nd, ppm | 1653.1237 | 113.7825 | 1425.5586 | 1880.6887 | 1311.7761 | 1994.4713 | 6.883% | 13.77% | 20.65% | 1570.4675 | 1735.7798
Pr, ppm | 483.9177 | 85.5539 | 312.81 | 655.0255 | 227.2561 | 740.5794 | 17.68% | 35.36% | 53.04% | 459.7219 | 508.1136
Th, ppm | 232.665 | 45.6019 | 141.4611 | 323.8689 | 95.8592 | 369.4708 | 19.6% | 39.2% | 58.8% | 221.0317 | 244.2982
Thermogravimetry
LOI, wt.% | 0.99639 | 0.175 | 0.6464 | 1.3464 | 0.4714 | 1.5214 | 17.56% | 35.13% | 52.69% | 0.94657 | 1.0462
Borate / Peroxide Fusion ICP
Al, wt.% | 5.9859 | 0.20778 | 5.5703 | 6.4015 | 5.3626 | 6.6093 | 3.471% | 6.942% | 10.41% | 5.6866 | 6.2852
Ba, ppm | 928.8402 | 55.2115 | 818.4172 | 1039.2633 | 763.2057 | 1094.4748 | 5.944% | 11.89% | 17.83% | 882.3982 | 975.2823
Be, ppm | 2.3019 | 0.376 | 1.5499 | 3.0539 | 1.1739 | 3.4299 | 16.33% | 32.67% | 49% | 2.1868 | 2.417
Bi, ppm | 2.4852 | 0.3867 | 1.7118 | 3.2586 | 1.3251 | 3.6453 | 15.56% | 31.12% | 46.68% | 2.3609 | 2.6095
Ca, wt.% | 1.2585 | 0.057298 | 1.1439 | 1.3731 | 1.0866 | 1.4304 | 4.553% | 9.106% | 13.66% | 1.1956 | 1.3214
Ce, ppm | 3509.7977 | 131.9597 | 3245.8782 | 3773.7172 | 3113.9185 | 3905.677 | 3.76% | 7.52% | 11.28% | 3334.3079 | 3685.2876
Co, ppm | 11.5257 | 1.7734 | 7.9789 | 15.0724 | 6.2055 | 16.8458 | 15.39% | 30.77% | 46.16% | 10.9494 | 12.102
Cr, ppm | 599.0322 | 40.4701 | 518.0921 | 679.9724 | 477.622 | 720.4425 | 6.756% | 13.51% | 20.27% | 569.0806 | 628.9838
Cs, ppm | 0.78598 | 0.058996 | 0.66799 | 0.90397 | 0.60899 | 0.96296 | 7.506% | 15.01% | 22.52% | 0.74668 | 0.82528
Cu, ppm | 60.1422 | 19.7586 | 20.6251 | 99.6593 | 0.86656 | 119.4179 | 32.85% | 65.71% | 98.56% | 57.1351 | 63.1493
Dy, ppm | 34.8324 | 1.444 | 31.9443 | 37.7204 | 30.5003 | 39.1644 | 4.146% | 8.291% | 12.44% | 33.0907 | 36.574
Er, ppm | 8.8026 | 0.29606 | 8.2105 | 9.3947 | 7.9144 | 9.6908 | 3.363% | 6.727% | 10.09% | 8.3625 | 9.2427
Eu, ppm | 46.7145 | 2.1891 | 42.3363 | 51.0927 | 40.1472 | 53.2818 | 4.686% | 9.372% | 14.06% | 44.3787 | 49.0502
Fe, wt.% | 31.7733 | 1.3012 | 29.1708 | 34.3758 | 27.8696 | 35.677 | 4.095% | 8.191% | 12.29% | 30.1847 | 33.362
Ga, ppm | 50.0389 | 5.8276 | 38.3836 | 61.6942 | 32.556 | 67.5218 | 11.65% | 23.29% | 34.94% | 47.5369 | 52.5408
Gd, ppm | 100.3873 | 8.068 | 84.2512 | 116.5233 | 76.1832 | 124.5913 | 8.037% | 16.07% | 24.11% | 95.3679 | 105.4066
Hf, ppm | 14.093 | 0.60698 | 12.879 | 15.3069 | 12.272 | 15.9139 | 4.307% | 8.614% | 12.92% | 13.3883 | 14.7976
Ho, ppm | 4.5595 | 0.30588 | 3.9477 | 5.1713 | 3.6419 | 5.4771 | 6.709% | 13.42% | 20.13% | 4.3315 | 4.7875
In, ppm | 0.60113 | 0.07596 | 0.44921 | 0.75305 | 0.37325 | 0.82901 | 12.64% | 25.27% | 37.91% | 0.57107 | 0.63118
K, wt.% | 0.23655 | 0.033137 | 0.17028 | 0.30282 | 0.13714 | 0.33596 | 14.01% | 28.02% | 42.02% | 0.22472 | 0.24838
La, ppm | 2690.0867 | 157.5096 | 2375.0675 | 3005.106 | 2217.5579 | 3162.6156 | 5.855% | 11.71% | 17.57% | 2555.5824 | 2824.5911
Li, ppm | 12.0055 | 1.5327 | 8.9402 | 15.0709 | 7.4075 | 16.6035 | 12.77% | 25.53% | 38.3% | 11.4052 | 12.6058
Lu, ppm | 0.5154 | 0.043771 | 0.42786 | 0.60294 | 0.38409 | 0.64671 | 8.493% | 16.99% | 25.48% | 0.48963 | 0.54117
Mg, wt.% | 1.047 | 0.051391 | 0.94421 | 1.1498 | 0.89282 | 1.2012 | 4.908% | 9.817% | 14.73% | 0.99465 | 1.0993
Mn, ppm | 652.6366 | 48.8555 | 554.9257 | 750.3475 | 506.0702 | 799.203 | 7.486% | 14.97% | 22.46% | 620.0048 | 685.2684
Mo, ppm | 45.6795 | 4.8335 | 36.0125 | 55.3465 | 31.179 | 60.18 | 10.58% | 21.16% | 31.74% | 43.3955 | 47.9635
Nb, ppm | 1296.3889 | 97.646 | 1101.0969 | 1491.681 | 1003.4508 | 1589.327 | 7.532% | 15.06% | 22.6% | 1231.5695 | 1361.2084
Nd, ppm | 1628.7803 | 54.2871 | 1520.2061 | 1737.3545 | 1465.919 | 1791.6416 | 3.333% | 6.666% | 9.999% | 1547.3413 | 1710.2193
Ni, ppm | 63.6945 | 5.4703 | 52.754 | 74.635 | 47.2837 | 80.1053 | 8.588% | 17.18% | 25.76% | 60.5098 | 66.8792
P, wt.% | 0.35872 | 0.019045 | 0.32063 | 0.39681 | 0.30158 | 0.41585 | 5.309% | 10.62% | 15.93% | 0.34078 | 0.37665
Pb, ppm | 106.2928 | 6.3097 | 93.6734 | 118.9122 | 87.3637 | 125.2219 | 5.936% | 11.87% | 17.81% | 100.9781 | 111.6074
Pr, ppm | 488.5896 | 20.9195 | 446.7506 | 530.4287 | 425.831 | 551.3482 | 4.282% | 8.563% | 12.84% | 464.1601 | 513.0191
Rb, ppm | 13.5154 | 0.5424 | 12.4306 | 14.6002 | 11.8882 | 15.1426 | 4.013% | 8.026% | 12.04% | 12.8397 | 14.1912
S, ppm | 517.6667 | 82.1688 | 353.329 | 682.0044 | 271.1601 | 764.1732 | 15.87% | 31.75% | 47.62% | 491.7833 | 543.55
Sb, ppm | 2.7613 | 0.21987 | 2.3216 | 3.2011 | 2.1017 | 3.4209 | 7.962% | 15.92% | 23.89% | 2.6233 | 2.8994
Si, wt.% | 14.7082 | 0.3598 | 13.9886 | 15.4277 | 13.6288 | 15.7875 | 2.446% | 4.892% | 7.339% | 13.9727 | 15.4436
Sm, ppm | 219.6468 | 7.3567 | 204.9334 | 234.3603 | 197.5767 | 241.717 | 3.349% | 6.699% | 10.05% | 208.6645 | 230.6292
Sn, ppm | 25.6303 | 3.3461 | 18.938 | 32.3226 | 15.5919 | 35.6687 | 13.06% | 26.11% | 39.17% | 24.3488 | 26.9118
Sr, ppm | 578.9394 | 15.3214 | 548.2966 | 609.5822 | 532.9752 | 624.9036 | 2.646% | 5.293% | 7.939% | 549.9924 | 607.8864
Ta, ppm | 25.1318 | 1.0401 | 23.0517 | 27.212 | 22.0116 | 28.2521 | 4.138% | 8.277% | 12.42% | 23.8752 | 26.3884
Tb, ppm | 9.0752 | 0.30531 | 8.4646 | 9.6858 | 8.1593 | 9.9912 | 3.364% | 6.728% | 10.09% | 8.6215 | 9.529
Th, ppm | 209.8552 | 6.8117 | 196.2319 | 223.4785 | 189.4202 | 230.2901 | 3.246% | 6.492% | 9.738% | 199.3624 | 220.3479
Ti, wt.% | 1.8399 | 0.055067 | 1.7298 | 1.95 | 1.6747 | 2.0051 | 2.993% | 5.986% | 8.979% | 1.7479 | 1.9319
Tm, ppm | 0.88945 | 0.085451 | 0.71855 | 1.0604 | 0.6331 | 1.1458 | 9.607% | 19.21% | 28.82% | 0.84498 | 0.93393
U, ppm | 4.7882 | 0.19497 | 4.3983 | 5.1782 | 4.2033 | 5.3732 | 4.072% | 8.144% | 12.22% | 4.5488 | 5.0277
V, ppm | 384.6194 | 16.4758 | 351.6679 | 417.5709 | 335.1921 | 434.0466 | 4.284% | 8.567% | 12.85% | 365.3884 | 403.8503
W, ppm | 3.8778 | 0.66844 | 2.5409 | 5.2147 | 1.8725 | 5.8831 | 17.24% | 34.48% | 51.71% | 3.6839 | 4.0717
Y, ppm | 90.9849 | 4.3596 | 82.2657 | 99.704 | 77.9061 | 104.0636 | 4.792% | 9.583% | 14.37% | 86.4356 | 95.5341
Yb, ppm | 4.3941 | 0.24773 | 3.8986 | 4.8895 | 3.6509 | 5.1372 | 5.638% | 11.28% | 16.91% | 4.1744 | 4.6138
Zn, ppm | 161.7333 | 23.352 | 115.0294 | 208.4373 | 91.6774 | 231.7893 | 14.44% | 28.88% | 43.32% | 153.6467 | 169.82
Zr, ppm | 603.0104 | 26.8975 | 549.2154 | 656.8055 | 522.3178 | 683.703 | 4.461% | 8.921% | 13.38% | 572.8599 | 633.1609
4-Acid Digestion
Ag, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Al, wt.% | 5.9585 | 0.18013 | 5.5983 | 6.3188 | 5.4181 | 6.4989 | 3.023% | 6.046% | 9.069% | 5.6606 | 6.2565
As, ppm | 33.5537 | 1.0854 | 31.3829 | 35.7245 | 30.2976 | 36.8099 | 3.235% | 6.47% | 9.704% | 31.876 | 35.2314
Ba, ppm | 957.0984 | 44.069 | 868.9604 | 1045.2365 | 824.8913 | 1089.3056 | 4.604% | 9.209% | 13.81% | 909.2435 | 1004.9534
Be, ppm | 2.0913 | 0.24371 | 1.6039 | 2.5787 | 1.3602 | 2.8224 | 11.65% | 23.31% | 34.96% | 1.9868 | 2.1959
Bi, ppm | 2.4269 | 0.15827 | 2.1104 | 2.7434 | 1.9521 | 2.9017 | 6.521% | 13.04% | 19.56% | 2.3055 | 2.5482
Ca, wt.% | 1.2147 | 0.090313 | 1.0341 | 1.3953 | 0.94376 | 1.4856 | 7.435% | 14.87% | 22.31% | 1.154 | 1.2754
Cd, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 3508.4834 | 111.9702 | 3284.543 | 3732.4237 | 3172.5728 | 3844.3939 | 3.191% | 6.383% | 9.574% | 3333.0592 | 3683.9075
Co, ppm | 12.1248 | 0.44431 | 11.2362 | 13.0134 | 10.7919 | 13.4577 | 3.665% | 7.329% | 10.99% | 11.5186 | 12.731
Cr, ppm | 523.6841 | 73.1099 | 377.4642 | 669.9039 | 304.3543 | 743.0139 | 13.96% | 27.92% | 41.88% | 497.4999 | 549.8683
Cs, ppm | 0.76887 | 0.066983 | 0.6349 | 0.90283 | 0.56792 | 0.96982 | 8.712% | 17.42% | 26.14% | 0.73042 | 0.80731
Cu, ppm | 55.8138 | 2.2885 | 51.2368 | 60.3909 | 48.9483 | 62.6794 | 4.1% | 8.201% | 12.3% | 53.0231 | 58.6045
Dy, ppm | 33.5055 | 0.86045 | 31.7847 | 35.2264 | 30.9242 | 36.0869 | 2.568% | 5.136% | 7.704% | 31.8303 | 35.1808
Er, ppm | 7.4597 | 0.25253 | 6.9546 | 7.9647 | 6.7021 | 8.2173 | 3.385% | 6.771% | 10.16% | 7.0867 | 7.8326
Eu, ppm | 47.6394 | 3.2615 | 41.1164 | 54.1624 | 37.8549 | 57.4239 | 6.846% | 13.69% | 20.54% | 45.2574 | 50.0214
Fe, wt.% | 31.5178 | 0.95936 | 29.5991 | 33.4365 | 28.6397 | 34.3959 | 3.044% | 6.088% | 9.132% | 29.9419 | 33.0937
Ga, ppm | 45.5202 | 8.0263 | 29.4677 | 61.5727 | 21.4415 | 69.599 | 17.63% | 35.26% | 52.9% | 43.2442 | 47.7963
Gd, ppm | 97.5977 | 3.4838 | 90.6301 | 104.5653 | 87.1464 | 108.049 | 3.57% | 7.139% | 10.71% | 92.7178 | 102.4776
Hf, ppm | 7.3912 | 0.83938 | 5.7124 | 9.0699 | 4.873 | 9.9093 | 11.36% | 22.71% | 34.07% | 7.0216 | 7.7607
Ho, ppm | 4.0576 | 0.11474 | 3.8281 | 4.2871 | 3.7134 | 4.4018 | 2.828% | 5.655% | 8.483% | 3.8547 | 4.2605
In, ppm | 0.56812 | 0.027058 | 0.51401 | 0.62224 | 0.48695 | 0.64929 | 4.763% | 9.525% | 14.29% | 0.53972 | 0.59653
K, wt.% | 0.23312 | 0.0059742 | 0.22118 | 0.24507 | 0.2152 | 0.25105 | 2.563% | 5.125% | 7.688% | 0.22147 | 0.24478
La, ppm | 2476.481 | 227.9521 | 2020.5769 | 2932.3852 | 1792.6248 | 3160.3373 | 9.205% | 18.41% | 27.61% | 2352.657 | 2600.3051
Li, ppm | 12.0397 | 1.3787 | 9.2824 | 14.7971 | 7.9037 | 16.1758 | 11.45% | 22.9% | 34.35% | 11.4378 | 12.6417
Lu, ppm | 0.38585 | 0.065651 | 0.25454 | 0.51715 | 0.18889 | 0.5828 | 17.01% | 34.03% | 51.04% | 0.36655 | 0.40514
Mg, wt.% | 0.99901 | 0.10801 | 0.783 | 1.215 | 0.675 | 1.323 | 10.81% | 21.62% | 32.43% | 0.94906 | 1.049
Mn, ppm | 547.2361 | 79.7765 | 387.6831 | 706.7891 | 307.9067 | 786.5656 | 14.58% | 29.16% | 43.73% | 519.8743 | 574.598
Mo, ppm | 48.2464 | 2.5361 | 43.1742 | 53.3187 | 40.6381 | 55.8548 | 5.257% | 10.51% | 15.77% | 45.8341 | 50.6588
Na, wt.% | 0.17817 | 0.01301 | 0.15215 | 0.2042 | 0.13914 | 0.21721 | 7.302% | 14.6% | 21.91% | 0.16927 | 0.18708
Nd, ppm | 1623.483 | 79.5193 | 1464.4444 | 1782.5217 | 1384.9251 | 1862.041 | 4.898% | 9.796% | 14.69% | 1542.3089 | 1704.6572
Ni, ppm | 71.1258 | 3.1439 | 64.838 | 77.4135 | 61.6942 | 80.5574 | 4.42% | 8.84% | 13.26% | 67.5695 | 74.6821
P, wt.% | 0.33785 | 0.02169 | 0.29447 | 0.38123 | 0.27278 | 0.40292 | 6.42% | 12.84% | 19.26% | 0.32096 | 0.35475
Pb, ppm | 103.9184 | 5.4831 | 92.9521 | 114.8846 | 87.469 | 120.3678 | 5.276% | 10.55% | 15.83% | 98.7225 | 109.1143
Pr, ppm | 482.2815 | 20.1118 | 442.0578 | 522.5052 | 421.946 | 542.6171 | 4.17% | 8.34% | 12.51% | 458.1675 | 506.3956
Rb, ppm | 13.2859 | 0.98994 | 11.306 | 15.2658 | 10.3161 | 16.2557 | 7.451% | 14.9% | 22.35% | 12.6216 | 13.9502
Re, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, ppm | 475.4111 | 48.6238 | 378.1635 | 572.6588 | 329.5396 | 621.2826 | 10.23% | 20.46% | 30.68% | 451.6406 | 499.1817
Sb, ppm | 2.4174 | 0.1736 | 2.0702 | 2.7646 | 1.8966 | 2.9382 | 7.181% | 14.36% | 21.54% | 2.2965 | 2.5383
Sc, ppm | 38.8492 | 2.505 | 33.8393 | 43.8591 | 31.3344 | 46.3641 | 6.448% | 12.9% | 19.34% | 36.9068 | 40.7917
Sm, ppm | 208.2366 | 7.4192 | 193.3981 | 223.075 | 185.9789 | 230.4943 | 3.563% | 7.126% | 10.69% | 197.8247 | 218.6484
Sn, ppm | 21.6377 | 2.0829 | 17.472 | 25.8034 | 15.3891 | 27.8862 | 9.626% | 19.25% | 28.88% | 20.5558 | 22.7196
Sr, ppm | 564.1131 | 35.1659 | 493.7813 | 634.445 | 458.6153 | 669.611 | 6.234% | 12.47% | 18.7% | 535.9075 | 592.3188
Ta, ppm | 20.9304 | 1.96 | 17.0103 | 24.8504 | 15.0503 | 26.8104 | 9.364% | 18.73% | 28.09% | 19.8838 | 21.9769
Tb, ppm | 9.0942 | 0.50289 | 8.0884 | 10.1 | 7.5855 | 10.6029 | 5.53% | 11.06% | 16.59% | 8.6395 | 9.5489
Te, ppm | 0.34127 | 0.055579 | 0.23011 | 0.45243 | 0.17453 | 0.50801 | 16.29% | 32.57% | 48.86% | 0.32421 | 0.35833
Th, ppm | 206.55 | 9.605 | 187.34 | 225.7599 | 177.7351 | 235.3649 | 4.65% | 9.3% | 13.95% | 196.2225 | 216.8775
Ti, wt.% | 1.0566 | 0.11708 | 0.82242 | 1.2907 | 0.70535 | 1.4078 | 11.08% | 22.16% | 33.24% | 1.0037 | 1.1094
Tl, ppm | 0.10384 | 0.0084602 | 0.086916 | 0.12076 | 0.078455 | 0.12922 | 8.148% | 16.3% | 24.44% | 0.098644 | 0.10903
Tm, ppm | 0.70061 | 0.046188 | 0.60824 | 0.79299 | 0.56205 | 0.83918 | 6.593% | 13.19% | 19.78% | 0.66558 | 0.73564
U, ppm | 4.4013 | 0.16104 | 4.0792 | 4.7233 | 3.9181 | 4.8844 | 3.659% | 7.318% | 10.98% | 4.1812 | 4.6213
V, ppm | 352.0608 | 28.763 | 294.5347 | 409.5868 | 265.7717 | 438.3498 | 8.17% | 16.34% | 24.51% | 334.4577 | 369.6638
W, ppm | 2.4971 | 0.36095 | 1.7752 | 3.219 | 1.4142 | 3.5799 | 14.45% | 28.91% | 43.36% | 2.3722 | 2.6219
Y, ppm | 83.097 | 3.7487 | 75.5996 | 90.5943 | 71.8509 | 94.343 | 4.511% | 9.022% | 13.53% | 78.9421 | 87.2518
Yb, ppm | 3.0631 | 0.14929 | 2.7645 | 3.3617 | 2.6153 | 3.511 | 4.874% | 9.748% | 14.62% | 2.91 | 3.2163
Zn, ppm | 139.3186 | 13.2856 | 112.7473 | 165.8898 | 99.4617 | 179.1754 | 9.536% | 19.07% | 28.61% | 132.3526 | 146.2845
Zr, ppm | 269.8554 | 18.0276 | 233.8002 | 305.9105 | 215.7726 | 323.9381 | 6.68% | 13.36% | 20.04% | 256.3626 | 283.3481

## Fusion XRF
| Analytical results for Al in OREAS 461 (Indicative Value 6.08 wt.%) | [8 empty] | Indicative
Al | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 6.092 | 5.927 | 6.192 | 6.245 |  |  |  | 1
| 1 | 2 | 6.086 | 5.98 | 6.086 | 6.139 |  |  |  | 1
| 1 | 3 | 6.065 | 5.98 | 6.086 | 6.192 |  |  |  | 16
| 1 | 4 | 6.102 | 5.822 | 6.139 | 6.192 |  |  |  | 6.0814
| 1 | 5 | 6.113 | 5.927 | 6.033 | 6.139 |  |  |  | 7
| 1 | 6 | 6.102 | 6.033 | 6.086 | 6.192
| Mean |  | 6.0933 | 5.9448 | 6.1037 | 6.1832
| Median |  | 6.097 | 5.9535 | 6.086 | 6.192
| Std Dev. |  | 0.016705 | 0.072071 | 0.054738 | 0.039897
| Rel.Std.Dev. |  | 0.2742% | 1.212% | 0.8968% | 0.6453%
| PDM3 |  | 0.196% | -2.246% | 0.3659% | 1.673%
| Z-Score (Absolute) |  | 0.08 | 2.3 | 0.08 | 1.27
| Analytical results for Ba in OREAS 461 (Indicative Value 1000 ppm) | [8 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 1070 | [5 empty] | 1
| 1 | 2 | N.A. | 990 | [5 empty] | 2
| 1 | 3 | N.A. | 990 | [5 empty] | 16
| 1 | 4 | N.A. | 900 | [5 empty] | 1000.1493
| 1 | 5 | N.A. | 1070 | [5 empty] | 8
| 1 | 6 | -- | 990
| Mean |  |  | 1001.6667
| Median |  |  | 990
| Std Dev. |  |  | 63.3772
| Rel.Std.Dev. |  |  | 6.327%
| PDM3 |  |  | 0.1517%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ca in OREAS 461 (Indicative Value 1.3 wt.%) | [8 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.286 | 1.322 | 1.258 | 1.344 |  |  |  | 1
| 1 | 2 | 1.286 | 1.301 | 1.265 | 1.336 |  |  |  | 3
| 1 | 3 | 1.286 | 1.286 | 1.251 | 1.344 |  |  |  | 16
| 1 | 4 | 1.286 | 1.294 | 1.265 | 1.344 |  |  |  | 1.2975
| 1 | 5 | 1.279 | 1.315 | 1.244 | 1.344 |  |  |  | 9
| 1 | 6 | 1.279 | 1.315 | 1.251 | 1.358
| Mean |  | 1.2837 | 1.3055 | 1.2557 | 1.345
| Median |  | 1.286 | 1.308 | 1.2545 | 1.344
| Std Dev. |  | 0.0036148 | 0.014039 | 0.0084774 | 0.0071274
| Rel.Std.Dev. |  | 0.2816% | 1.075% | 0.6751% | 0.5299%
| PDM3 |  | -1.064% | 0.6191% | -3.222% | 3.663%
| Z-Score (Absolute) |  | 0.29 | 0.29 | 1.06 | 1.35
| Analytical results for Ce in OREAS 461 (Certified Value 3562 ppm) | [8 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 3660 | 3580 | 3590 | 3580 | 3760 | 3600 | 3500 | 1
| 1 | 2 | 3580 | 3580 | 3500 | 3500 | 3760 | 3600 | 3500 | #N/A
| 1 | 3 | 3580 | 3580 | 3330 | 3500 | 3760 | 3600 | 3580 | 16
| 1 | 4 | 3580 | 3580 | 3590 | 3580 | 3670 | 3600 | 3500 | 3562.359
| 1 | 5 | 3580 | 3500 | 3590 | 3500 | 3760 | 3600 | 3500 | 7
| 1 | 6 | 3580 | 3580 | 3590 | 3500 | 3760 | 3600 | 3500
| Mean |  | 3593.3333 | 3566.6667 | 3531.6667 | 3526.6667 | 3745 | 3600 | 3513.3333
| Median |  | 3580 | 3580 | 3590 | 3500 | 3760 | 3600 | 3500
| Std Dev. |  | 32.6599 | 32.6599 | 105.1507 | 41.3118 | 36.7423 | 0 | 32.6599
| Rel.Std.Dev. |  | 0.9089% | 0.9157% | 2.977% | 1.171% | 0.9811% | 0% | 0.9296%
| PDM3 |  | 0.8695% | 0.1209% | -0.8616% | -1.002% | 5.127% | 1.057% | -1.376%
| Z-Score (Absolute) |  | 0.58 | 0.01 | 0 | 0.88 | 3.74 | 0.67 | 1.18
| Analytical results for Cr in OREAS 461 (Indicative Value 430 ppm) | [8 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 404 | [5 empty] | 1
| 1 | 2 | N.A. | 424 | [5 empty] | 4
| 1 | 3 | N.A. | 424 | [5 empty] | 16
| 1 | 4 | N.A. | 452 | [5 empty] | 429.9236
| 1 | 5 | N.A. | 438 | [5 empty] | 10
| 1 | 6 | -- | 438
| Mean |  |  | 430
| Median |  |  | 431
| Std Dev. |  |  | 16.4924
| Rel.Std.Dev. |  |  | 3.835%
| PDM3 |  |  | 0.01777%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Dy in OREAS 461 (Indicative Value 40.7 ppm) | [8 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | 170 | N.A. | 44 |  |  |  |  | 1
| 1 | 2 | 170 | N.A. | 35 |  |  |  |  | 5
| 1 | 3 | 170 | N.A. | 44 |  |  |  |  | 16
| 1 | 4 | 170 | N.A. | 44 |  |  |  |  | 40.661
| 1 | 5 | 170 | N.A. | 35 |  |  |  |  | 11
| 1 | 6 | 170 | -- | 44
| Mean |  | 170 |  | 41
| Median |  | 170 |  | 44
| Std Dev. |  | 3.1134e-14 |  | 4.6476
| Rel.Std.Dev. |  | 1.831e-14% |  | 11.34%
| PDM3 |  | 318.1% |  | 0.8336%
| Z-Score (Absolute) |  | 0.67 | NA | 0.67
| Analytical results for Er in OREAS 461 (Indicative Value 17.5 ppm) | [8 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 90 | N.A. | 17 |  |  |  |  | 1
| 1 | 2 | < 90 | N.A. | < 9 |  |  |  |  | 6
| 1 | 3 | < 90 | N.A. | 17 |  |  |  |  | 16
| 1 | 4 | < 90 | N.A. | 17 |  |  |  |  | 17.4902
| 1 | 5 | < 90 | N.A. | 17 |  |  |  |  | 12
| 1 | 6 | < 90 | -- | 17
| Mean |  |  |  | 17
| Median |  |  |  | 17
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  |  | -2.802%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Eu in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 09 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 90 | < 90 | < 90 |  |  |  |  | 1
| 1 | 2 | < 90 | < 90 | < 90 |  |  |  |  | 7
| 1 | 3 | < 90 | < 90 | < 90 |  |  |  |  | 16
| 1 | 4 | < 90 | < 90 | < 90 |  |  |  |  | < 100
| 1 | 5 | < 90 | < 90 | < 90 |  |  |  |  | 13
| 1 | 6 | < 90 | < 90 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Fe in OREAS 461 (Certified Value 32.24 wt.%) | [8 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 09 | 13 | 14 | 15 | 22 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 32.245 | 33.154 | 31.68 | 31.685 | 32.45 |  |  | 1
| 1 | 2 | 32.433 | 32.874 | 31.62 | 31.615 | 32.52 |  |  | #N/A
| 1 | 3 | 32.531 | 32.734 | 31.75 | 31.615 | 32.45 |  |  | 16
| 1 | 4 | 32.524 | 32.804 | 31.68 | 31.545 | 32.38 |  |  | 32.2378
| 1 | 5 | 32.398 | 33.084 | 31.55 | 31.615 | 32.38 |  |  | 8
| 1 | 6 | 32.461 | 33.364 | 31.68 | 31.825 | 32.45
| Mean |  | 32.432 | 33.0023 | 31.66 | 31.65 | 32.4383
| Median |  | 32.447 | 32.979 | 31.68 | 31.615 | 32.45
| Std Dev. |  | 0.1051 | 0.24012 | 0.067823 | 0.096488 | 0.052694
| Rel.Std.Dev. |  | 0.3241% | 0.7276% | 0.2142% | 0.3049% | 0.1624%
| PDM3 |  | 0.6025% | 2.372% | -1.792% | -1.823% | 0.6222%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.91 | 0.93 | 0.01
| Analytical results for Gd in OREAS 461 (Indicative Value 87 ppm) | [8 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 09 | 14 | 22 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 90 | 90 | < 90 | 90 |  |  |  | 1
| 1 | 2 | 90 | 170 | < 90 | 90 |  |  |  | 8
| 1 | 3 | 90 | 90 | < 90 | 90 |  |  |  | 16
| 1 | 4 | 90 | 90 | 90 | 90 |  |  |  | 86.7604
| 1 | 5 | 90 | 170 | < 90 | 90 |  |  |  | 14
| 1 | 6 | 90 | 90 | 90 | 90
| Mean |  | 90 | 116.6667 | 90 | 90
| Median |  | 90 | 90 | 90 | 90
| Std Dev. |  | 0 | 41.3118 | 0 | 0
| Rel.Std.Dev. |  | 0% | 35.41% | 0% | 0%
| PDM3 |  | 3.734% | 34.47% | 3.734% | 3.734%
| Z-Score (Absolute) |  | 0 | 1.35 | 1.35 | 0
| Analytical results for Hf in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Hf | Round | Replicate | Lab | [6 empty] | 1
| No | No | 13 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 9
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 15
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 461 (Indicative Value 26.2 ppm) | [8 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | 90 | 9 | [5 empty] | 1
| 1 | 2 | < 90 | < 9 | [5 empty] | 10
| 1 | 3 | < 90 | 9 | [5 empty] | 16
| 1 | 4 | < 90 | 9 | [5 empty] | 26.1894
| 1 | 5 | < 90 | 9 | [5 empty] | 16
| 1 | 6 | < 90 | 9
| Mean |  | 90 | 9
| Median |  | 90 | 9
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  | 243.7% | -65.64%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for K in OREAS 461 (Indicative Value 0.236 wt.%) | [8 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.232 | 0.332 | 0.232 | 0.241 |  |  |  | 1
| 1 | 2 | 0.232 | 0.315 | 0.232 | 0.241 |  |  |  | 11
| 1 | 3 | 0.232 | 0.299 | 0.232 | 0.241 |  |  |  | 16
| 1 | 4 | 0.232 | 0.315 | 0.241 | 0.241 |  |  |  | 0.23576
| 1 | 5 | 0.232 | 0.315 | 0.232 | 0.241 |  |  |  | 17
| 1 | 6 | 0.232 | 0.324 | 0.224 | 0.241
| Mean |  | 0.232 | 0.31667 | 0.23217 | 0.241
| Median |  | 0.232 | 0.315 | 0.232 | 0.241
| Std Dev. |  | 0 | 0.011039 | 0.0053821 | 3.0405e-17
| Rel.Std.Dev. |  | 0% | 3.486% | 2.318% | 1.262e-14%
| PDM3 |  | -1.596% | 34.32% | -1.525% | 2.221%
| Z-Score (Absolute) |  | 0.81 | 12.9 | 0.54 | 0.54
| Analytical results for La in OREAS 461 (Certified Value 2720 ppm) | [8 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 2640 | 2730 | 2810 | 2730 | 2730 | 2800 | 2730 | 1
| 1 | 2 | 2640 | 2730 | 2730 | 2730 | 2730 | 2800 | 2640 | #N/A
| 1 | 3 | 2730 | 2730 | 2730 | 2640 | 2640 | 2800 | 2730 | 16
| 1 | 4 | 2640 | 2730 | 2640 | 2560 | 2730 | 2800 | 2730 | 2720.4541
| 1 | 5 | 2730 | 2730 | 2730 | 2730 | 2730 | 2800 | 2730 | 9
| 1 | 6 | 2730 | 2730 | 2640 | 2730 | 2730 | 2800 | 2730
| Mean |  | 2685 | 2730 | 2713.3333 | 2686.6667 | 2715 | 2800 | 2715
| Median |  | 2685 | 2730 | 2730 | 2730 | 2730 | 2800 | 2730
| Std Dev. |  | 49.295 | 0 | 64.7045 | 71.7403 | 36.7423 | 0 | 36.7423
| Rel.Std.Dev. |  | 1.836% | 0% | 2.385% | 2.67% | 1.353% | 0% | 1.353%
| PDM3 |  | -1.303% | 0.3509% | -0.2617% | -1.242% | -0.2005% | 2.924% | -0.2005%
| Z-Score (Absolute) |  | 1.35 | 0.67 | 0 | 1.35 | 0 | 4.07 | 0
| Analytical results for Lu in OREAS 461 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 18 | [5 empty] | 1
| 1 | 2 | < 90 | < 18 | [5 empty] | 1
| 1 | 3 | < 90 | < 18 | [5 empty] | 16
| 1 | 4 | < 90 | < 18 | [5 empty] | < 20
| 1 | 5 | < 90 | < 18 | [5 empty] | 7
| 1 | 6 | < 90 | < 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Mg in OREAS 461 (Indicative Value 1.06 wt.%) | [8 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.079 | 1.019 | 1.073 | 1.067 |  |  |  | 1
| 1 | 2 | 1.086 | 1.037 | 1.079 | 1.067 |  |  |  | 2
| 1 | 3 | 1.098 | 1.013 | 1.067 | 1.079 |  |  |  | 16
| 1 | 4 | 1.092 | 1.031 | 1.061 | 1.043 |  |  |  | 1.0611
| 1 | 5 | 1.079 | 1.013 | 1.073 | 1.061 |  |  |  | 8
| 1 | 6 | 1.086 | 1.025 | 1.049 | 1.086
| Mean |  | 1.0867 | 1.023 | 1.067 | 1.0672
| Median |  | 1.086 | 1.022 | 1.07 | 1.067
| Std Dev. |  | 0.0074207 | 0.009798 | 0.010733 | 0.014945
| Rel.Std.Dev. |  | 0.6829% | 0.9578% | 1.006% | 1.4%
| PDM3 |  | 2.406% | -3.594% | 0.5522% | 0.5679%
| Z-Score (Absolute) |  | 1.35 | 3.12 | 0 | 0
| Analytical results for Mn in OREAS 461 (Indicative Value 700 ppm) | [8 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 700 | 770 | 620 | 770 |  |  |  | 1
| 1 | 2 | 700 | 700 | 620 | 770 |  |  |  | 3
| 1 | 3 | 700 | 770 | 620 | 700 |  |  |  | 16
| 1 | 4 | 700 | 700 | 620 | 700 |  |  |  | 700.253
| 1 | 5 | 700 | 700 | 620 | 770 |  |  |  | 9
| 1 | 6 | 700 | 770 | 620 | 770
| Mean |  | 700 | 735 | 620 | 746.6667
| Median |  | 700 | 735 | 620 | 770
| Std Dev. |  | 0 | 38.3406 | 0 | 36.1478
| Rel.Std.Dev. |  | 0% | 5.216% | 0% | 4.841%
| PDM3 |  | -0.03613% | 4.962% | -11.46% | 6.628%
| Z-Score (Absolute) |  | 0.51 | 0.51 | 2.53 | 0.84
| Analytical results for Na in OREAS 461 (Indicative Value 0.186 wt.%) | [8 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.193 | 0.185 | 0.178 | 0.2 |  |  |  | 1
| 1 | 2 | 0.171 | 0.178 | 0.23 | 0.178 |  |  |  | 4
| 1 | 3 | 0.193 | 0.193 | 0.208 | 0.178 |  |  |  | 16
| 1 | 4 | 0.171 | 0.2 | 0.185 | 0.185 |  |  |  | 0.18558
| 1 | 5 | 0.178 | 0.171 | 0.178 | 0.185 |  |  |  | 10
| 1 | 6 | 0.193 | 0.178 | 0.193 | 0.193
| Mean |  | 0.18317 | 0.18417 | 0.19533 | 0.1865
| Median |  | 0.1855 | 0.1815 | 0.189 | 0.185
| Std Dev. |  | 0.011071 | 0.010759 | 0.020373 | 0.0086429
| Rel.Std.Dev. |  | 6.044% | 5.842% | 10.43% | 4.634%
| PDM3 |  | -1.302% | -0.7635% | 5.254% | 0.4938%
| Z-Score (Absolute) |  | 0.9 | 0.45 | 1.08 | 0.45
| Analytical results for Nb in OREAS 461 (Indicative Value 1302 ppm) | [8 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | N.A. | 1260 | 1300 |  |  |  |  | 1
| 1 | 2 | N.A. | 1330 | 1300 |  |  |  |  | 5
| 1 | 3 | N.A. | 1260 | 1300 |  |  |  |  | 16
| 1 | 4 | N.A. | 1330 | 1300 |  |  |  |  | 1302.4525
| 1 | 5 | N.A. | 1330 | 1300 |  |  |  |  | 11
| 1 | 6 | -- | 1330 | 1300
| Mean |  |  | 1306.6667 | 1300
| Median |  |  | 1330 | 1300
| Std Dev. |  |  | 36.1478 | 0
| Rel.Std.Dev. |  |  | 2.766% | 0%
| PDM3 |  |  | 0.3236% | -0.1883%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Nd in OREAS 461 (Certified Value 1653 ppm) | [8 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 1710 | 1540 | 1370 | 1710 | 1800 | 1700 | 1603 | 1
| 1 | 2 | 1710 | 1630 | 1630 | 1540 | 1890 | 1700 | 1612 | #N/A
| 1 | 3 | 1710 | 1630 | 1460 | 1710 | 1890 | 1700 | 1629 | 16
| 1 | 4 | 1710 | 1630 | 1460 | 1710 | 1800 | 1700 | 1629 | 1653.1237
| 1 | 5 | 1710 | 1630 | 1460 | 1540 | 1800 | 1700 | 1620 | 10
| 1 | 6 | 1710 | 1630 | 1460 | 1630 | 1800 | 1700 | 1629
| Mean |  | 1710 | 1615 | 1473.3333 | 1640 | 1830 | 1700 | 1620.3333
| Median |  | 1710 | 1630 | 1460 | 1670 | 1800 | 1700 | 1624.5
| Std Dev. |  | 0 | 36.7423 | 84.7742 | 83.4266 | 46.4758 | 0 | 10.9118
| Rel.Std.Dev. |  | 0% | 2.275% | 5.754% | 5.087% | 2.54% | 0% | 0.6734%
| PDM3 |  | 3.441% | -2.306% | -10.88% | -0.7939% | 10.7% | 2.836% | -1.984%
| Z-Score (Absolute) |  | 0.79 | 0.17 | 2.41 | 0 | 2.21 | 0.67 | 0.27
| Analytical results for P in OREAS 461 (Indicative Value 0.371 wt.%) | [8 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 09 | 14 | 15 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 0.3666 | 0.3692 | 0.375 |  |  |  |  | 1
| 1 | 2 | 0.3657 | 0.3666 | 0.375 |  |  |  |  | 6
| 1 | 3 | 0.3648 | 0.3722 | 0.375 |  |  |  |  | 16
| 1 | 4 | 0.3657 | 0.37 | 0.375 |  |  |  |  | 0.37099
| 1 | 5 | 0.3679 | 0.3722 | 0.38 |  |  |  |  | 12
| 1 | 6 | 0.3644 | 0.3722 | 0.38
| Mean |  | 0.36585 | 0.3704 | 0.37667
| Median |  | 0.3657 | 0.3711 | 0.375
| Std Dev. |  | 0.0012661 | 0.0022698 | 0.002582
| Rel.Std.Dev. |  | 0.3461% | 0.6128% | 0.6855%
| PDM3 |  | -1.386% | -0.1597% | 1.529%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.93
| Analytical results for Pr in OREAS 461 (Certified Value 484 ppm) | [8 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 510 | 500 | 430 | 330 | 510 | 400 | 580 | 1
| 1 | 2 | 510 | 500 | 940 | 330 | 510 | 500 | 500 | #N/A
| 1 | 3 | 510 | 500 | 850 | 330 | 600 | 500 | 580 | 16
| 1 | 4 | 510 | 410 | 1030 | 330 | 510 | 500 | 580 | 483.9177
| 1 | 5 | 430 | 500 | 430 | 330 | 430 | 500 | 500 | 11
| 1 | 6 | 510 | 500 | 680 | 330 | 680 | 500 | 580
| Mean |  | 496.6667 | 485 | 726.6667 | 330 | 540 | 483.3333 | 553.3333
| Median |  | 510 | 500 | 765 | 330 | 510 | 500 | 580
| Std Dev. |  | 32.6599 | 36.7423 | 257.2677 | 0 | 87.178 | 40.8248 | 41.3118
| Rel.Std.Dev. |  | 6.576% | 7.576% | 35.4% | 0% | 16.14% | 8.447% | 7.466%
| PDM3 |  | 2.635% | 0.2236% | 50.16% | -31.81% | 11.59% | -0.1208% | 14.34%
| Z-Score (Absolute) |  | 0.05 | 0.28 | 3.71 | 2.77 | 0.67 | 0 | 0.85
| Analytical results for Si in OREAS 461 (Indicative Value 14.8 wt.%) | [8 empty] | Indicative
Si | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 14.8 | 15.052 | 14.772 | 14.772 |  |  |  | 1
| 1 | 2 | 14.805 | 14.912 | 14.819 | 14.725 |  |  |  | 7
| 1 | 3 | 14.809 | 14.725 | 14.819 | 14.772 |  |  |  | 16
| 1 | 4 | 14.828 | 14.491 | 15.052 | 14.678 |  |  |  | 14.7966
| 1 | 5 | 14.739 | 14.819 | 14.725 | 14.725 |  |  |  | 13
| 1 | 6 | 14.73 | 14.959 | 14.865 | 14.725
| Mean |  | 14.7852 | 14.8263 | 14.842 | 14.7328
| Median |  | 14.8025 | 14.8655 | 14.819 | 14.725
| Std Dev. |  | 0.040474 | 0.19932 | 0.11338 | 0.03538
| Rel.Std.Dev. |  | 0.2737% | 1.344% | 0.7639% | 0.2401%
| PDM3 |  | -0.07736% | 0.2009% | 0.3067% | -0.431%
| Z-Score (Absolute) |  | 0.49 | 0.49 | 0.86 | 1.73
| Analytical results for Sm in OREAS 461 (Indicative Value 190 ppm) | [8 empty] | Indicative
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 14 | 15 | 16 | 22 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | < 90 | 260 | 170 | < 430 | 200 | 170 |  | 1
| 1 | 2 | < 90 | 260 | 170 | < 430 | 200 | 170 |  | 8
| 1 | 3 | 90 | 260 | 170 | < 430 | 200 | 170 |  | 16
| 1 | 4 | 90 | 170 | 170 | < 430 | 200 | 170 |  | 190.1345
| 1 | 5 | 90 | 170 | 90 | < 430 | 200 | 170 |  | 14
| 1 | 6 | 90 | 170 | 170 | < 430 | 200 | 170
| Mean |  | 90 | 215 | 156.6667 |  | 200 | 170
| Median |  | 90 | 215 | 170 |  | 200 | 170
| Std Dev. |  | 0 | 49.295 | 32.6599 |  | 3.1134e-14 | 0
| Rel.Std.Dev. |  | 0% | 22.93% | 20.85% |  | 1.557e-14% | 0%
| PDM3 |  | -52.67% | 13.08% | -17.6% |  | 5.189% | -10.59%
| Z-Score (Absolute) |  | 2.39 | 1.06 | 0 | NA | 0.67 | 0
| Analytical results for Sn in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | < 80 | [5 empty] | 1
| 1 | 2 | N.A. | < 80 | [5 empty] | 9
| 1 | 3 | N.A. | < 80 | [5 empty] | 16
| 1 | 4 | N.A. | < 80 | [5 empty] | < 100
| 1 | 5 | N.A. | < 80 | [5 empty] | 15
| 1 | 6 | -- | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for S in OREAS 461 (Indicative Value 0.047 wt.%) | [8 empty] | Indicative
S | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.0485 | [6 empty] | 1
| 1 | 2 | 0.0444 | [6 empty] | 10
| 1 | 3 | 0.0461 | [6 empty] | 16
| 1 | 4 | 0.044 | [6 empty] | 0.046786
| 1 | 5 | 0.0473 | [6 empty] | 16
| 1 | 6 | 0.0505
| Mean |  | 0.0468
| Median |  | 0.0467
| Std Dev. |  | 0.0024868
| Rel.Std.Dev. |  | 5.314%
| PDM3 |  | 0.03049%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 461 (Indicative Value 423 ppm) | [8 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 420 | [5 empty] | 1
| 1 | 2 | N.A. | 420 | [5 empty] | 11
| 1 | 3 | N.A. | 510 | [5 empty] | 16
| 1 | 4 | N.A. | 340 | [5 empty] | 422.7972
| 1 | 5 | N.A. | 420 | [5 empty] | 17
| 1 | 6 | -- | 420
| Mean |  |  | 421.6667
| Median |  |  | 420
| Std Dev. |  |  | 53.8207
| Rel.Std.Dev. |  |  | 12.76%
| PDM3 |  |  | -0.2674%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ta in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | < 80 | [5 empty] | 1
| 1 | 2 | N.A. | < 80 | [5 empty] | 1
| 1 | 3 | N.A. | < 80 | [5 empty] | 16
| 1 | 4 | N.A. | < 80 | [5 empty] | < 100
| 1 | 5 | N.A. | < 80 | [5 empty] | 7
| 1 | 6 | -- | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Tb in OREAS 461 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 17 | [5 empty] | 1
| 1 | 2 | < 90 | < 17 | [5 empty] | 2
| 1 | 3 | < 90 | < 17 | [5 empty] | 16
| 1 | 4 | < 90 | < 17 | [5 empty] | < 20
| 1 | 5 | < 90 | < 17 | [5 empty] | 8
| 1 | 6 | < 90 | < 17
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 461 (Certified Value 233 ppm) | [8 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 180 | 260 | < 260 | 260 | 281 | 170 |  | 1
| 1 | 2 | 190 | 260 | < 260 | 260 | 290 | 170 |  | #N/A
| 1 | 3 | 200 | 260 | 350 | 260 | 316 | 160 |  | 16
| 1 | 4 | 180 | 260 | < 260 | 260 | 281 | 170 |  | 232.665
| 1 | 5 | 200 | 260 | < 260 | 260 | 272 | 170 |  | 12
| 1 | 6 | 190 | 260 | < 260 | 260 | 264 | 170
| Mean |  | 190 | 260 | 350 | 260 | 284 | 168.3333
| Median |  | 190 | 260 | 350 | 260 | 281 | 170
| Std Dev. |  | 8.9443 | 0 |  | 0 | 18.0111 | 4.0825
| Rel.Std.Dev. |  | 4.708% | 0% |  | 0% | 6.342% | 2.425%
| PDM3 |  | -18.34% | 11.75% | 50.43% | 11.75% | 22.06% | -27.65%
| Z-Score (Absolute) |  | 0.57 | 0.57 | 1.46 | 0.57 | 0.78 | 0.9
| Analytical results for Ti in OREAS 461 (Indicative Value 1.84 wt.%) | [8 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.846 | 1.852 | 1.804 | 1.9 |  |  |  | 1
| 1 | 2 | 1.834 | 1.834 | 1.816 | 1.894 |  |  |  | 3
| 1 | 3 | 1.84 | 1.816 | 1.81 | 1.9 |  |  |  | 16
| 1 | 4 | 1.84 | 1.798 | 1.81 | 1.894 |  |  |  | 1.8439
| 1 | 5 | 1.834 | 1.834 | 1.792 | 1.894 |  |  |  | 9
| 1 | 6 | 1.834 | 1.84 | 1.828 | 1.9
| Mean |  | 1.838 | 1.829 | 1.81 | 1.897
| Median |  | 1.837 | 1.834 | 1.81 | 1.897
| Std Dev. |  | 0.004899 | 0.019131 | 0.012 | 0.0032863
| Rel.Std.Dev. |  | 0.2665% | 1.046% | 0.663% | 0.1732%
| PDM3 |  | -0.3207% | -0.8088% | -1.839% | 2.879%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 1.13 | 3.06
| Analytical results for Tm in OREAS 461 (Indicative Value < 10 ppm) | [8 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 9 | [5 empty] | 1
| 1 | 2 | < 90 | < 9 | [5 empty] | 4
| 1 | 3 | < 90 | < 9 | [5 empty] | 16
| 1 | 4 | < 90 | < 9 | [5 empty] | < 10
| 1 | 5 | < 90 | < 9 | [5 empty] | 10
| 1 | 6 | < 90 | < 9
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 80 | < 250 | N.A. | 80 |  |  |  | 1
| 1 | 2 | < 80 | < 250 | N.A. | < 80 |  |  |  | 5
| 1 | 3 | < 80 | < 250 | N.A. | < 80 |  |  |  | 16
| 1 | 4 | < 80 | < 250 | N.A. | < 80 |  |  |  | < 100
| 1 | 5 | < 80 | < 250 | N.A. | < 80 |  |  |  | 11
| 1 | 6 | < 80 | < 250 | -- | < 80
| Mean |  |  |  |  | 80
| Median |  |  |  |  | 80
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for V in OREAS 461 (Indicative Value 408 ppm) | [8 empty] | Indicative
V | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 415 | [5 empty] | 1
| 1 | 2 | N.A. | 398 | [5 empty] | 6
| 1 | 3 | N.A. | 415 | [5 empty] | 16
| 1 | 4 | N.A. | 403 | [5 empty] | 407.9842
| 1 | 5 | N.A. | 415 | [5 empty] | 12
| 1 | 6 | -- | 403
| Mean |  |  | 408.1667
| Median |  |  | 409
| Std Dev. |  |  | 7.705
| Rel.Std.Dev. |  |  | 1.888%
| PDM3 |  |  | 0.04473%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for W in OREAS 461 (Indicative Value < 100 ppm) | [8 empty] | Indicative
W | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 7
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 13
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 461 (Indicative Value 115 ppm) | [8 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | < 100 | < 80 | < 80 | 160 | 160 | < 100 | 94 | 1
| 1 | 2 | < 100 | 80 | < 80 | 160 | 80 | < 100 | 94 | 8
| 1 | 3 | < 100 | < 80 | < 80 | 160 | 80 | < 100 | 94 | 16
| 1 | 4 | < 100 | 80 | < 80 | 160 | < 80 | < 100 | 71 | 115.1665
| 1 | 5 | < 100 | 160 | < 80 | 80 | 160 | < 100 | 71 | 14
| 1 | 6 | < 100 | 160 | < 80 | 160 | 240 | < 100 | 94
| Mean |  |  | 120 |  | 146.6667 | 144 |  | 86.3333
| Median |  |  | 120 |  | 160 | 160 |  | 94
| Std Dev. |  |  | 46.188 |  | 32.6599 | 66.9328 |  | 11.8771
| Rel.Std.Dev. |  |  | 38.49% |  | 22.27% | 46.48% |  | 13.76%
| PDM3 |  |  | 4.197% |  | 27.35% | 25.04% |  | -25.04%
| Z-Score (Absolute) |  | NA | 0.58 | NA | 1.74 | 0.58 | NA | 0.77
| Analytical results for Yb in OREAS 461 (Indicative Value 8.78 ppm) | [8 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 2
2
| 1 | 1 | < 90 | N.A. | 9 |  |  |  |  | 1
| 1 | 2 | < 90 | N.A. | < 9 |  |  |  |  | 9
| 1 | 3 | < 90 | N.A. | 9 |  |  |  |  | 16
| 1 | 4 | < 90 | N.A. | 9 |  |  |  |  | 8.7819
| 1 | 5 | < 90 | N.A. | 9 |  |  |  |  | 15
| 1 | 6 | < 90 | -- | 9
| Mean |  |  |  | 9
| Median |  |  |  | 9
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  |  | 2.483%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Zr in OREAS 461 (Indicative Value 592 ppm) | [8 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 590 | [5 empty] | 1
| 1 | 2 | N.A. | 590 | [5 empty] | 10
| 1 | 3 | N.A. | 590 | [5 empty] | 16
| 1 | 4 | N.A. | 590 | [5 empty] | 592.2416
| 1 | 5 | N.A. | 590 | [5 empty] | 16
| 1 | 6 | -- | 590
| Mean |  |  | 590
| Median |  |  | 590
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | -0.3785%
| Z-Score (Absolute) |  | NA | NA
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 461 (Indicative Value 0.697 wt.%) | [8 empty] | Indicative
H2O- | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | at 1000°C | [6 empty] | 3
3
| 1 | 1 | 0.68 | [6 empty] | 1
| 1 | 2 | 0.73 | [6 empty] | 13
| 1 | 3 | 0.71 | [6 empty] | 16
| 1 | 4 | 0.7 | [6 empty] | 0.69667
| 1 | 5 | 0.69 | [6 empty] | 19
| 1 | 6 | 0.67
| Mean |  | 0.69667
| Median |  | 0.695
| Std Dev. |  | 0.021602
| Rel.Std.Dev. |  | 3.101%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI in OREAS 461 (Certified Value 0.996 wt.%) | [8 empty] | Certified
LOI | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 18 | wt.%
|  |  | at 1000°C | at 1000°C | at 1000°C | BF*OES | at 1000°C | BF*XRF | at 1000°C | 3
3
| 1 | 1 | 0.92 | 0.84 | 0.99 | 1.12 | 1.3 | 0.776 | 1.74 | 1
| 1 | 2 | 0.93 | 0.8 | 1.01 | 1.13 | 1.3 | 0.832 | 1.7793 | #N/A
| 1 | 3 | 0.94 | 0.83 | 0.99 | 1.13 | 1.28 | 0.83 | 1.7379 | 16
| 1 | 4 | 0.94 | 0.9 | 0.96 | 1.01 | 1.31 | 0.823 | 1.6371 | 0.99639
| 1 | 5 | 0.98 | 0.82 | 0.96 | 1.2 | 1.28 | 0.812 | 1.679 | 14
| 1 | 6 | 0.94 | 0.83 | 0.97 | 1.12 | 1.25 | 0.763 | 1.6793
| Mean |  | 0.94167 | 0.83667 | 0.98 | 1.1183 | 1.2867 | 0.806 | 1.7088
| Median |  | 0.94 | 0.83 | 0.98 | 1.125 | 1.29 | 0.8175 | 1.7086
| Std Dev. |  | 0.020412 | 0.033862 | 0.02 | 0.061128 | 0.021602 | 0.029414 | 0.052334
| Rel.Std.Dev. |  | 2.168% | 4.047% | 2.041% | 5.466% | 1.679% | 3.649% | 3.063%
| PDM3 |  | -5.492% | -16.03% | -1.645% | 12.24% | 29.13% | -19.11% | 71.5%
| Z-Score (Absolute) |  | 0.16 | 0.66 | 0 | 0.67 | 1.29 | 0.73 | 3.07
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 461 (Indicative Value 15.2 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 20 | 22 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 5 | 18 | < 5 | [12 empty] | 1
| 1 | 2 | < 5 | 14 | < 5 | [12 empty] | 15
| 1 | 3 | < 5 | 14 | < 5 | [12 empty] | 16
| 1 | 4 | < 5 | 14 | < 5 | [12 empty] | 15.1667
| 1 | 5 | < 5 | 16 | < 5 | [12 empty] | 21
| 1 | 6 | < 5 | 15 | < 5
| Mean |  |  | 15.1667
| Median |  |  | 14.5
| Std Dev. |  |  | 1.6021
| Rel.Std.Dev. |  |  | 10.56%
| PDM3 |  |  | -2.22e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Al in OREAS 461 (Certified Value 5.99 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 6.044 | 6.21 | 5.96 | 5.9853 | 6.002 | 5.6855 | 5.71 | 6.05 | 6.23 | 6.15 | [5 empty] | 1
| 1 | 2 | 6.061 | 6.17 | 5.98 | 6.0567 | 6.16 | 5.6521 | 5.58 | 5.88 | 6.28 | 6.14 | [5 empty] | 1
| 1 | 3 | 6.043 | 6.11 | 5.95 | 6.1052 | 6.039 | 5.6969 | 5.56 | 5.72 | 6.14 | 5.83 | [5 empty] | 16
| 1 | 4 | 5.998 | 6.19 | 5.98 | 5.9231 | 6.155 | 5.7252 | 5.64 | 5.8 | 6.21 | 6.2 | [5 empty] | 5.9859
| 1 | 5 | 6.007 | 6.14 | 6.03 | 6.0802 | 6.049 | 5.6742 | 5.51 | 5.9 | 6.33 | 6.13 | [5 empty] | 16
| 1 | 6 | 5.963 | 6.01 | 5.96 | 6.0116 | 5.996 | 5.7591 | 5.61 | 6.06 | 6.41 | 6.19
| Mean |  | 6.0193 | 6.1383 | 5.9767 | 6.027 | 6.0668 | 5.6988 | 5.6017 | 5.9017 | 6.2667 | 6.1067
| Median |  | 6.025 | 6.155 | 5.97 | 6.0341 | 6.044 | 5.6912 | 5.595 | 5.89 | 6.255 | 6.145
| Std Dev. |  | 0.036566 | 0.072226 | 0.028752 | 0.067248 | 0.073167 | 0.038213 | 0.069113 | 0.13482 | 0.095219 | 0.13837
| Rel.Std.Dev. |  | 0.6075% | 1.177% | 0.4811% | 1.116% | 1.206% | 0.6705% | 1.234% | 2.284% | 1.519% | 2.266%
| PDM3 |  | 0.5585% | 2.546% | -0.1543% | 0.6868% | 1.352% | -4.796% | -6.419% | -1.407% | 4.69% | 2.017%
| Z-Score (Absolute) |  | 0.02 | 0.66 | 0.27 | 0.02 | 0.25 | 1.85 | 2.4 | 0.69 | 1.39 | 0.79
| Analytical results for As in OREAS 461 (Indicative Value 79 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 37 | < 100 | 79.8798 | 100 | 40 | 140 | [9 empty] | 1
| 1 | 2 | 39 | < 100 | 80.4216 | 100 | 50 | 150 | [9 empty] | 16
| 1 | 3 | 38 | < 100 | 78.2879 | 90 | 40 | 130 | [9 empty] | 16
| 1 | 4 | 35 | < 100 | 83.1789 | 80 | 40 | 140 | [9 empty] | 79.3256
| 1 | 5 | 37 | < 100 | 80.0767 | 110 | 40 | 150 | [9 empty] | 22
| 1 | 6 | 36 | < 100 | 79.474 | 90 | 30 | 160
| Mean |  | 37 |  | 80.2198 | 95 | 40 | 145
| Median |  | 37 |  | 79.9782 | 95 | 40 | 145
| Std Dev. |  | 1.4142 |  | 1.6262 | 10.4881 | 6.3246 | 10.4881
| Rel.Std.Dev. |  | 3.822% |  | 2.027% | 11.04% | 15.81% | 7.233%
| PDM3 |  | -53.36% |  | 1.127% | 19.76% | -49.57% | 82.79%
| Z-Score (Absolute) |  | 0.73 | NA | 0 | 0.26 | 0.67 | 1.11
| Analytical results for B in OREAS 461 (Indicative Value 151 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*OES | PF*OES | [13 empty] | 0
0
| 1 | 1 | 158 | 140 | [13 empty] | 1
| 1 | 2 | 157 | 140 | [13 empty] | 17
| 1 | 3 | 165 | 140 | [13 empty] | 16
| 1 | 4 | 161 | 160 | [13 empty] | 150.8333
| 1 | 5 | 176 | 140 | [13 empty] | 23
| 1 | 6 | 153 | 120
| Mean |  | 161.6667 | 140
| Median |  | 159.5 | 140
| Std Dev. |  | 8.0911 | 12.6491
| Rel.Std.Dev. |  | 5.005% | 9.035%
| PDM3 |  | 7.182% | -7.182%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ba in OREAS 461 (Certified Value 929 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 977 | 901 | 979 | 972 | 927 | 853.7079 | 963 | 922.7 | 820.264 | 914 | 950 | 985.9762 | 930 | 990 | 800 | 1
| 1 | 2 | 1010 | 950 | 971 | 944 | 911 | 857.1178 | 953 | 921 | 825.016 | 934 | 950 | 992.5471 | 920 | 970 | 820 | 2
| 1 | 3 | 1020 | 905 | 1000 | 972 | 937 | 869.7307 | 944 | 904.7 | 826.152 | 905 | 950 | 954.9139 | 930 | 970 | 920 | 16
| 1 | 4 | 998 | 883 | 969 | 984 | 952 | 859.5518 | 922 | 916.5 | 824.416 | 932 | 950 | 977.1231 | 920 | 970 | 840 | 928.8402
| 1 | 5 | 992 | 945 | 1005 | 988 | 935 | 871.9238 | 979 | 899.9 | 820.424 | 930 | 950 | 952.2215 | 890 | 1010 | 810 | 17
| 1 | 6 | 976 | 911 | 975 | 990 | 946 | 891.7386 | 923 | 916.7 | 819.584 | 912 | 950 | 969.7136 | 930 | 1020 | 830
| Mean |  | 995.5 | 915.8333 | 983.1667 | 975 | 934.6667 | 867.2951 | 947.3333 | 913.5833 | 822.6427 | 921.1667 | 950 | 972.0825 | 920 | 988.3333 | 836.6667
| Median |  | 995 | 908 | 977 | 978 | 936 | 864.6412 | 948.5 | 916.6 | 822.42 | 922 | 950 | 973.4183 | 925 | 980 | 825
| Std Dev. |  | 17.6153 | 26.2939 | 15.4456 | 17.0529 | 14.5144 | 13.9498 | 22.4737 | 9.1923 | 2.8646 | 12.3031 | 1.2454e-13 | 16.3222 | 15.4919 | 22.286 | 43.2049
| Rel.Std.Dev. |  | 1.769% | 2.871% | 1.571% | 1.749% | 1.553% | 1.608% | 2.372% | 1.006% | 0.3482% | 1.336% | 1.311e-14% | 1.679% | 1.684% | 2.255% | 5.164%
| PDM3 |  | 7.177% | -1.4% | 5.849% | 4.97% | 0.6273% | -6.626% | 1.991% | -1.643% | -11.43% | -0.8261% | 2.278% | 4.656% | -0.9518% | 6.405% | -9.924%
| Z-Score (Absolute) |  | 1.1 | 0.34 | 0.87 | 0.73 | 0 | 1.21 | 0.23 | 0.38 | 2.02 | 0.24 | 0.28 | 0.67 | 0.16 | 0.97 | 2.07
| Analytical results for Be in OREAS 461 (Certified Value 2.3 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 17 | 22 | 23 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 2.4177 | 3 | 2.2 | 1.1375 | 2 | 2 | < 5 | [8 empty] | 1
| 1 | 2 | 2.1938 | 3 | 2.2 | 0.97849 | 2 | 2 | < 5 | [8 empty] | #N/A
| 1 | 3 | 2.2676 | 3 | 2.3 | 1.0965 | 2 | 2 | < 5 | [8 empty] | 16
| 1 | 4 | 2.2226 | 2 | 2.3 | 1.18 | 2 | 2 | < 5 | [8 empty] | 2.3019
| 1 | 5 | 2.1672 | 3 | 2.3 | 1.0649 | 2 | 3 | < 5 | [8 empty] | 18
| 1 | 6 | 2.2874 | 3 | 2.2 | 1.1239 | 2 | 2 | < 5
| Mean |  | 2.2594 | 2.8333 | 2.25 | 1.0969 | 2 | 2.1667
| Median |  | 2.2451 | 3 | 2.25 | 1.1102 | 2 | 2
| Std Dev. |  | 0.08949 | 0.40825 | 0.054772 | 0.069761 | 0 | 0.40825
| Rel.Std.Dev. |  | 3.961% | 14.41% | 2.434% | 6.36% | 0% | 18.84%
| PDM3 |  | -1.846% | 23.09% | -2.254% | -52.35% | -13.11% | -5.874%
| Z-Score (Absolute) |  | 0.27 | 3.25 | 0.22 | 5.78 | 1.08 | 0.22 | NA
| Analytical results for Bi in OREAS 461 (Certified Value 2.49 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 2.6178 | 2.6 | 2 | 1.2 | 3 | 2.4 | [9 empty] | 1
| 1 | 2 | 2.6915 | 2.6 | 2 | 1.1 | 2 | 2.2 | [9 empty] | #N/A
| 1 | 3 | 3.154 | 2.6 | 2 | 1.1 | 3 | 2.4 | [9 empty] | 16
| 1 | 4 | 3.0122 | 2.5 | 2 | 1.3 | 2 | 2.3 | [9 empty] | 2.4852
| 1 | 5 | 2.7663 | 2.6 | 2 | 1.1 | 3 | 2.1 | [9 empty] | 19
| 1 | 6 | 2.9142 | 2.6 | 2.5 | 1.2 | 3 | 2
| Mean |  | 2.8594 | 2.5833 | 2.0833 | 1.1667 | 2.6667 | 2.2333
| Median |  | 2.8403 | 2.6 | 2 | 1.15 | 3 | 2.25
| Std Dev. |  | 0.20416 | 0.040825 | 0.20412 | 0.08165 | 0.5164 | 0.1633
| Rel.Std.Dev. |  | 7.14% | 1.58% | 9.798% | 6.999% | 19.36% | 7.312%
| PDM3 |  | 15.06% | 3.949% | -16.17% | -53.06% | 7.302% | -10.13%
| Z-Score (Absolute) |  | 1.04 | 0.4 | 0.75 | 2.87 | 0.6 | 0.4
| Analytical results for Ca in OREAS 461 (Certified Value 1.26 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.27 | 1.3 | 1.29 | 1.2643 | 1.215 | 1.34 | 1.2473 | 1.294 | 1.286 | 1.4 | 1.2 |  |  |  |  | 1
| 1 | 2 | 1.263 | 1.3 | 1.29 | 1.1514 | 1.351 | 1.23 | 1.2509 | 1.315 | 1.233 | 1.4 | 1.2 |  |  |  |  | 3
| 1 | 3 | 1.242 | 1.3 | 1.28 | 1.1462 | 1.229 | 1.27 | 1.2552 | 1.286 | 1.189 | 1.3 | 1.2 |  |  |  |  | 16
| 1 | 4 | 1.257 | 1.3 | 1.28 | 1.1535 | 1.236 | 1.24 | 1.2609 | 1.301 | 1.206 | 1.3 | 1.2 |  |  |  |  | 1.2585
| 1 | 5 | 1.261 | 1.4 | 1.28 | 1.1648 | 1.215 | 1.25 | 1.2352 | 1.265 | 1.235 | 1.4 | 1.2 |  |  |  |  | 20
| 1 | 6 | 1.238 | 1.3 | 1.28 | 1.19 | 1.208 | 1.3 | 1.2545 | 1.301 | 1.293 | 1.4 | 1.2
| Mean |  | 1.2552 | 1.3167 | 1.2833 | 1.1784 | 1.2423 | 1.2717 | 1.2507 | 1.2937 | 1.2403 | 1.3667 | 1.2
| Median |  | 1.259 | 1.3 | 1.28 | 1.1591 | 1.222 | 1.26 | 1.2527 | 1.2975 | 1.234 | 1.4 | 1.2
| Std Dev. |  | 0.012545 | 0.040825 | 0.005164 | 0.044894 | 0.054221 | 0.041673 | 0.0088435 | 0.01699 | 0.041846 | 0.05164 | 0
| Rel.Std.Dev. |  | 0.9994% | 3.101% | 0.4024% | 3.81% | 4.364% | 3.277% | 0.7071% | 1.313% | 3.374% | 3.779% | 0%
| PDM3 |  | -0.2634% | 4.623% | 1.975% | -6.367% | -1.283% | 1.048% | -0.6218% | 2.796% | -1.442% | 8.597% | -4.647%
| Z-Score (Absolute) |  | 0 | 0.88 | 0.55 | 1.84 | 0.67 | 0.32 | 0.09 | 0.75 | 0.29 | 2.18 | 1.08
| Analytical results for Cd in OREAS 461 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | < 1 | < 1 | < 10 | < 0.2 | [11 empty] | 1
| 1 | 2 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | 18
| 1 | 3 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | 16
| 1 | 4 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | < 1
| 1 | 5 | < 1 | < 1 | < 10 | 0.2 | [11 empty] | 24
| 1 | 6 | < 1 | < 1 | < 10 | 0.4
| Mean |  |  |  |  | 0.3
| Median |  |  |  |  | 0.3
| Std Dev. |  |  |  |  | 0.070711
| Rel.Std.Dev. |  |  |  |  | 23.57%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Ce in OREAS 461 (Certified Value 3510 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 3640 | 3380 | 3570 | 3520 | 3370 | 3293.4359 | 3667 | 3237.9 | 3297.9101 | 2834.6 | 3542 | 4038.0393 | 3640 | 3580 | 3450 | 1
| 1 | 2 | 3660 | 3680 | 3540 | 3420 | 3420 | 3283.6501 | 3645 | 3409 | 3258.0689 | 2811 | 3577 | 4060.1307 | 3500 | 3600 | 3450 | #N/A
| 1 | 3 | 3740 | 3450 | 3580 | 3540 | 3490 | 3338.4525 | 3623 | 3460 | 3367.7068 | 2764.6 | 3618 | 3864.7996 | 3560 | 3570 | 3770 | 16
| 1 | 4 | 3630 | 3370 | 3520 | 3570 | 3490 | 3323.9957 | 3611 | 3328.2 | 3299.1879 | 2724.7 | 3593 | 3944.638 | 3650 | 3610 | 3460 | 3509.7977
| 1 | 5 | 3610 | 3640 | 3600 | 3540 | 3460 | 3327.309 | 3607 | 3370 | 3361.5273 | 2781.7 | 3594 | 3915.9415 | 3710 | 3650 | 3470 | 21
| 1 | 6 | 3620 | 3530 | 3530 | 3560 | 3470 | 3365.047 | 3581 | 3228.7 | 3337.1329 | 2710.8 | 3723 | 3968.9306 | 3680 | 3750 | 3550
| Mean |  | 3650 | 3508.3333 | 3556.6667 | 3525 | 3450 | 3321.9817 | 3622.3333 | 3338.9667 | 3320.2556 | 2771.2333 | 3607.8333 | 3965.4133 | 3623.3333 | 3626.6667 | 3525
| Median |  | 3635 | 3490 | 3555 | 3540 | 3465 | 3325.6523 | 3617 | 3349.1 | 3318.1604 | 2773.15 | 3593.5 | 3956.7843 | 3645 | 3605 | 3465
| Std Dev. |  | 47.3286 | 131.3646 | 31.4113 | 54.3139 | 46.9042 | 29.8078 | 30.2699 | 92.7426 | 42.5445 | 48.0995 | 61.746 | 73.8383 | 78.6554 | 66.5332 | 125.8173
| Rel.Std.Dev. |  | 1.297% | 3.744% | 0.8832% | 1.541% | 1.36% | 0.8973% | 0.8356% | 2.778% | 1.281% | 1.736% | 1.711% | 1.862% | 2.171% | 1.835% | 3.569%
| PDM3 |  | 3.995% | -0.04172% | 1.335% | 0.4331% | -1.704% | -5.351% | 3.206% | -4.867% | -5.4% | -21.04% | 2.793% | 12.98% | 3.235% | 3.33% | 0.4331%
| Z-Score (Absolute) |  | 0.86 | 0.11 | 0.22 | 0 | 0.51 | 1.39 | 0.67 | 1.28 | 1.4 | 5.17 | 0.57 | 3.02 | 0.67 | 0.7 | 0.34
| Analytical results for Co in OREAS 461 (Certified Value 11.5 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 13.5 | 13.9258 | 9.9 | < 50 | 9.5234 | 9 | < 10 | 20 | 11.6 | [6 empty] | 1
| 1 | 2 | 13.5 | 13.8322 | 9.7 | < 50 | 9.4401 | 11.1 | < 10 | 20 | 10.9 | [6 empty] | #N/A
| 1 | 3 | 13.7 | 13.831 | 9.7 | < 50 | 9.4924 | 10.8 | < 10 | 20 | 13.8 | [6 empty] | 16
| 1 | 4 | 13.3 | 14.0516 | 9.9 | < 50 | 10.8107 | 6.9 | < 10 | 20 | 10.6 | [6 empty] | 11.5257
| 1 | 5 | 13.3 | 13.7931 | 9.9 | < 50 | 9.2972 | 11.5 | < 10 | 20 | 11.5 | [6 empty] | 22
| 1 | 6 | 13.1 | 13.8425 | 9.9 | < 50 | 9.7244 | 10.9 | < 10 | 20 | 11.6
| Mean |  | 13.4 | 13.8794 | 9.8333 |  | 9.7147 | 10.0333 |  | 20 | 11.6667
| Median |  | 13.4 | 13.8373 | 9.9 |  | 9.5079 | 10.85 |  | 20 | 11.55
| Std Dev. |  | 0.20976 | 0.095082 | 0.10328 |  | 0.55443 | 1.7614 |  | 0 | 1.1237
| Rel.Std.Dev. |  | 1.565% | 0.6851% | 1.05% |  | 5.707% | 17.56% |  | 0% | 9.632%
| PDM3 |  | 16.26% | 20.42% | -14.68% |  | -15.71% | -12.95% |  | 73.53% | 1.223%
| Z-Score (Absolute) |  | 0.64 | 0.81 | 0.67 | NA | 0.72 | 0.37 | NA | 3.07 | 0
| Analytical results for Cr in OREAS 461 (Certified Value 599 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 0
0
| 1 | 1 | 650 | 570 | 620 | 670 | 550 | < 1000 | 580 | 600.592 | 620 | 600 | 565.8703 | 600 | 750 | 450 | 610 | 1
| 1 | 2 | 660 | 610 | 610 | 660 | 540 | < 1000 | 571 | 598.927 | 620 | 600 | 585.7028 | 600 | 700 | 450 | 610 | 4
| 1 | 3 | 670 | 580 | 600 | 670 | 560 | < 1000 | 562 | 592.78 | 620 | 600 | 575.1069 | 500 | 770 | 470 | 660 | 16
| 1 | 4 | 660 | 570 | 620 | 670 | 550 | < 1000 | 574 | 594.518 | 620 | 600 | 575.2096 | 500 | 770 | 450 | 600 | 599.0322
| 1 | 5 | 650 | 610 | 620 | 670 | 540 | < 1000 | 567 | 593.529 | 620 | 600 | 566.7535 | 500 | 580 | 450 | 610 | 23
| 1 | 6 | 650 | 580 | 620 | 680 | 550 | < 1000 | 559 | 596.346 | 620 | 600 | 563.1512 | 600 | 730 | 470 | 600
| Mean |  | 656.6667 | 586.6667 | 615 | 670 | 548.3333 |  | 568.8333 | 596.1153 | 620 | 600 | 571.9657 | 550 | 716.6667 | 456.6667 | 615
| Median |  | 655 | 580 | 620 | 670 | 550 |  | 569 | 595.432 | 620 | 600 | 570.9302 | 550 | 740 | 450 | 610
| Std Dev. |  | 8.165 | 18.619 | 8.3666 | 6.3246 | 7.5277 |  | 7.7825 | 3.1101 | 0 | 1.2454e-13 | 8.3799 | 54.7723 | 72.0185 | 10.328 | 22.5832
| Rel.Std.Dev. |  | 1.243% | 3.174% | 1.36% | 0.944% | 1.373% |  | 1.368% | 0.5217% | 0% | 2.076e-14% | 1.465% | 9.959% | 10.05% | 2.262% | 3.672%
| PDM3 |  | 9.621% | -2.064% | 2.666% | 11.85% | -8.463% |  | -5.041% | -0.4869% | 3.5% | 0.1616% | -4.518% | -8.185% | 19.64% | -23.77% | 2.666%
| Z-Score (Absolute) |  | 1.43 | 0.28 | 0.49 | 1.75 | 1.21 | NA | 0.71 | 0.05 | 0.43 | 0.05 | 0.64 | 1.17 | 3.56 | 3.45 | 0.19
| Analytical results for Cs in OREAS 461 (Certified Value 0.79 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.87 | 0.76 | 0.83 | 0.77 | 0.74 | 0.77175 | 0.9 | 0.8 | 0.57 | < 1 | 0.57329 | 0.3 | 1 | 0.8 |  | 1
| 1 | 2 | 0.84 | 0.8 | 0.8 | 0.76 | 0.72 | 0.80135 | 0.8 | 0.8 | 0.55 | < 1 | 0.54669 | 0.3 | < 1 | 0.9 |  | #N/A
| 1 | 3 | 0.9 | 0.59 | 0.76 | 0.77 | 0.74 | 0.8125 | 0.8 | 0.8 | 0.52 | < 1 | 0.51027 | 0.3 | < 1 | 0.9 |  | 16
| 1 | 4 | 0.79 | 0.76 | 0.79 | 0.75 | 0.75 | 0.78179 | 0.7 | 0.8 | 0.53 | < 1 | 0.54631 | 0.2 | < 1 | 1.1 |  | 0.78598
| 1 | 5 | 0.79 | 0.71 | 0.82 | 0.69 | 0.79 | 0.77057 | 0.7 | 0.8 | 0.52 | < 1 | 0.52435 | 0.3 | < 1 | 1.1 |  | 24
| 1 | 6 | 0.8 | 0.64 | 0.79 | 0.76 | 0.73 | 0.77685 | 0.8 | 0.8 | 0.5 | < 1 | 0.5113 | 0.3 | < 1 | 1.1
| Mean |  | 0.83167 | 0.71 | 0.79833 | 0.75 | 0.745 | 0.7858 | 0.78333 | 0.8 | 0.53167 |  | 0.53537 | 0.28333 | 1 | 0.98333
| Median |  | 0.82 | 0.735 | 0.795 | 0.76 | 0.74 | 0.77932 | 0.8 | 0.8 | 0.525 |  | 0.53533 | 0.3 | 1 | 1
| Std Dev. |  | 0.046224 | 0.080498 | 0.024833 | 0.030332 | 0.02429 | 0.017205 | 0.075277 | 1.2162e-16 | 0.024833 |  | 0.024563 | 0.040825 |  | 0.13292
| Rel.Std.Dev. |  | 5.558% | 11.34% | 3.111% | 4.044% | 3.26% | 2.19% | 9.61% | 1.52e-14% | 4.671% |  | 4.588% | 14.41% |  | 13.52%
| PDM3 |  | 5.813% | -9.667% | 1.572% | -4.577% | -5.214% | -0.02273% | -0.3364% | 1.784% | -32.36% |  | -31.89% | -63.95% | 27.23% | 25.11%
| Z-Score (Absolute) |  | 0.67 | 1.02 | 0.21 | 0.3 | 0.66 | 0.03 | 0 | 0.23 | 3.51 | NA | 3.46 | 6.98 | 3.02 | 1.16
| Analytical results for Cu in OREAS 461 (Certified Value 60 ppm) | [16 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | 58 | < 100 | 69.7187 | 66 | < 80 | 50 | 60 | [8 empty] | 1
| 1 | 2 | 58 | < 100 | 71.7051 | 476 | < 80 | 50 | 60 | [8 empty] | #N/A
| 1 | 3 | 57 | < 100 | 76.0173 | 19 | < 80 | 100 | 60 | [8 empty] | 16
| 1 | 4 | 70 | < 100 | 70.4566 | 29 | < 80 | 50 | 60 | [8 empty] | 60.1422
| 1 | 5 | 58 | < 100 | 68.527 | 24 | < 80 | 100 | 60 | [8 empty] | 25
| 1 | 6 | 73 | < 100 | 69.0418 | 35 | < 80 | 100 | 60
| Mean |  | 62.3333 |  | 70.9111 | 108.1667 |  | 75 | 60
| Median |  | 58 |  | 70.0877 | 32 |  | 75 | 60
| Std Dev. |  | 7.174 |  | 2.739 | 180.9612 |  | 27.3861 | 0
| Rel.Std.Dev. |  | 11.51% |  | 3.863% | 167.3% |  | 36.51% | 0%
| PDM3 |  | 3.643% |  | 17.91% | 79.85% |  | 24.7% | -0.2365%
| Z-Score (Absolute) |  | 0 | NA | 0.67 | 1.61 | NA | 0.93 | 0.01
| Analytical results for Dy in OREAS 461 (Certified Value 34.8 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 32.4 | 34.4 | 34.1 | 35.1 | 32.9 | 35.4043 | 37.2 | 34.9 | 29.9536 | 28.03 | 35 | 33.5857 | 57.2 | 42.5 | 33.2 | 1
| 1 | 2 | 32.4 | 37.4 | 35.3 | 34.3 | 32.9 | 34.7576 | 36.8 | 34.9 | 30.2719 | 28.28 | 35 | 34.0409 | 54.6 | 35 | 33.7 | 5
| 1 | 3 | 33.8 | 36 | 34.3 | 35 | 33.5 | 36.4326 | 36.1 | 34.6 | 27.0546 | 27.89 | 36 | 32.4442 | 54.5 | 36 | 36.8 | 16
| 1 | 4 | 32.7 | 33.3 | 35.8 | 35.9 | 34.3 | 36.0649 | 35.5 | 34.8 | 28.2474 | 27.47 | 36 | 33.3189 | 47.1 | 38 | 35.1 | 34.8324
| 1 | 5 | 33 | 36.9 | 35.1 | 35.5 | 34 | 35.5006 | 36.9 | 34 | 30.0643 | 27.71 | 35 | 31.9604 | 58.7 | 37 | 33.8 | 26
| 1 | 6 | 32.5 | 34.1 | 33.5 | 35.3 | 34.2 | 35.3212 | 35.8 | 34.6 | 29.442 | 27.33 | 37 | 32.3987 | 55.9 | 37.5 | 35.1
| Mean |  | 32.8 | 35.35 | 34.6833 | 35.1833 | 33.6333 | 35.5802 | 36.3833 | 34.6333 | 29.1723 | 27.785 | 35.6667 | 32.9581 | 54.6667 | 37.6667 | 34.6167
| Median |  | 32.6 | 35.2 | 34.7 | 35.2 | 33.75 | 35.4525 | 36.45 | 34.7 | 29.6978 | 27.8 | 35.5 | 32.8816 | 55.25 | 37.25 | 34.45
| Std Dev. |  | 0.54037 | 1.655 | 0.85888 | 0.53821 | 0.6314 | 0.59034 | 0.67946 | 0.33862 | 1.2671 | 0.35439 | 0.8165 | 0.80857 | 4.0372 | 2.6013 | 1.3227
| Rel.Std.Dev. |  | 1.647% | 4.682% | 2.476% | 1.53% | 1.877% | 1.659% | 1.868% | 0.9777% | 4.344% | 1.275% | 2.289% | 2.453% | 7.385% | 6.906% | 3.821%
| PDM3 |  | -5.835% | 1.486% | -0.4278% | 1.008% | -3.442% | 2.147% | 4.453% | -0.5714% | -16.25% | -20.23% | 2.395% | -5.381% | 56.94% | 8.137% | -0.6192%
| Z-Score (Absolute) |  | 1.34 | 0.43 | 0 | 0.32 | 0.67 | 0.58 | 1.09 | 0.03 | 3.54 | 4.43 | 0.63 | 1.11 | 13.81 | 1.3 | 0.04
| Analytical results for Er in OREAS 461 (Certified Value 8.8 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 8.33 | 8.89 | 8.87 | 8.7 | 8.45 | 8.8171 | 9.1 | 8.7 | 8.639 | 7.17 | 8.8 | 8.0703 | 7.04 | 10 | 9.05 | 1
| 1 | 2 | 8.23 | 9.23 | 9.04 | 9.04 | 8.75 | 9.2253 | 9.3 | 8.7 | 8.3708 | 7 | 8.9 | 8.0932 | 6.69 | 9.5 | 9.14 | 6
| 1 | 3 | 8.73 | 9.18 | 9.18 | 8.66 | 8.89 | 8.8902 | 9.2 | 8.6 | 8.4003 | 6.81 | 9 | 7.8924 | 6.65 | 9.5 | 8.89 | 16
| 1 | 4 | 8.58 | 8.83 | 9.33 | 8.76 | 9.01 | 8.4512 | 8.9 | 8.5 | 8.6529 | 6.97 | 8.7 | 7.9835 | 5.79 | 9.5 | 9.01 | 8.8026
| 1 | 5 | 8.33 | 9.09 | 9.05 | 8.99 | 8.97 | 8.254 | 9.3 | 8.5 | 8.4507 | 6.91 | 8.5 | 7.724 | 7.1 | 9.5 | 9.57 | 27
| 1 | 6 | 8.14 | 8.91 | 8.76 | 9.11 | 8.75 | 8.716 | 8.8 | 8.5 | 8.2749 | 6.83 | 9 | 7.6831 | 6.82 | 10 | 8.95
| Mean |  | 8.39 | 9.0217 | 9.0383 | 8.8767 | 8.8033 | 8.7256 | 9.1 | 8.5833 | 8.4648 | 6.9483 | 8.8167 | 7.9078 | 6.6817 | 9.6667 | 9.1017
| Median |  | 8.33 | 9 | 9.045 | 8.875 | 8.82 | 8.7665 | 9.15 | 8.55 | 8.4255 | 6.94 | 8.85 | 7.938 | 6.755 | 9.5 | 9.03
| Std Dev. |  | 0.22226 | 0.16714 | 0.20547 | 0.19273 | 0.20422 | 0.34146 | 0.20976 | 0.098319 | 0.15165 | 0.13182 | 0.19408 | 0.17377 | 0.47301 | 0.2582 | 0.24482
| Rel.Std.Dev. |  | 2.649% | 1.853% | 2.273% | 2.171% | 2.32% | 3.913% | 2.305% | 1.145% | 1.792% | 1.897% | 2.201% | 2.197% | 7.079% | 2.671% | 2.69%
| PDM3 |  | -4.687% | 2.489% | 2.678% | 0.8416% | 0.008547% | -0.8741% | 3.379% | -2.491% | -3.838% | -21.06% | 0.16% | -10.17% | -24.09% | 9.816% | 3.398%
| Z-Score (Absolute) |  | 1.19 | 0.63 | 0.67 | 0.21 | 0 | 0.22 | 0.85 | 0.63 | 0.97 | 5.32 | 0.04 | 2.57 | 5.58 | 2.48 | 0.59
| Analytical results for Eu in OREAS 461 (Certified Value 46.7 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 43.9 | 46.9 | 44.8 | 46.4 | 43.7 | 44.5498 | 51.2 | 44.1 | 28.7559 | 39.04 | 44.5 | 48.9103 | 88 | 53 | 46.6 | 1
| 1 | 2 | 44.1 | 49.6 | 45.7 | 44.6 | 45.3 | 46.0219 | 49.8 | 44.8 | 28.8895 | 38.73 | 45.5 | 49.5536 | 83.3 | 47.5 | 47 | 7
| 1 | 3 | 45.9 | 48.3 | 45.4 | 45.2 | 45.5 | 45.1982 | 50.6 | 44.6 | 28.2768 | 38.34 | 46 | 47.1482 | 83.1 | 49.5 | 52.6 | 16
| 1 | 4 | 44.8 | 45.1 | 46.7 | 46.8 | 46.1 | 44.6746 | 50.5 | 45.4 | 29.5162 | 37.91 | 46 | 48.3393 | 71.8 | 48.5 | 45 | 46.7145
| 1 | 5 | 44.1 | 49.3 | 46.4 | 46.1 | 46.2 | 45.0665 | 51.2 | 44.1 | 29.5656 | 38.5 | 46 | 47.4818 | 88.2 | 50.5 | 47.4 | 28
| 1 | 6 | 44.1 | 45.8 | 45.9 | 46.9 | 46 | 45.7589 | 49.7 | 43.9 | 30.1127 | 38.07 | 47.5 | 48.9387 | 85.4 | 51 | 48
| Mean |  | 44.4833 | 47.5 | 45.8167 | 46 | 45.4667 | 45.2117 | 50.5 | 44.4833 | 29.1861 | 38.4317 | 45.9167 | 48.3953 | 83.3 | 50 | 47.7667
| Median |  | 44.1 | 47.6 | 45.8 | 46.25 | 45.75 | 45.1324 | 50.55 | 44.35 | 29.2028 | 38.42 | 46 | 48.6248 | 84.35 | 50 | 47.2
| Std Dev. |  | 0.76004 | 1.8601 | 0.68532 | 0.91869 | 0.93524 | 0.58375 | 0.65115 | 0.56362 | 0.66505 | 0.41835 | 0.9704 | 0.92684 | 6.0465 | 1.9494 | 2.575
| Rel.Std.Dev. |  | 1.709% | 3.916% | 1.496% | 1.997% | 2.057% | 1.291% | 1.289% | 1.267% | 2.279% | 1.089% | 2.113% | 1.915% | 7.259% | 3.899% | 5.391%
| PDM3 |  | -4.776% | 1.682% | -1.922% | -1.529% | -2.671% | -3.217% | 8.104% | -4.776% | -37.52% | -17.73% | -1.708% | 3.598% | 78.32% | 7.033% | 2.252%
| Z-Score (Absolute) |  | 0.67 | 0.74 | 0.05 | 0.04 | 0.21 | 0.33 | 2.16 | 0.67 | 7.87 | 3.52 | 0 | 1.17 | 18.67 | 1.92 | 0.42
| Analytical results for Fe in OREAS 461 (Certified Value 31.77 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 30.2971 | 33.46 | 31.67 | 32.6106 | 30.174 | 33.04 | 31.8269 | 30.286 | 31.92 | 33 | > 25 |  |  |  |  | 1
| 1 | 2 | 30.4225 | 33.47 | 31.89 | 33.5958 | 30.755 | 32.51 | 31.759 | 29.517 | 31.43 | 33.1 | > 25 |  |  |  |  | #N/A
| 1 | 3 | 30.3646 | 32.34 | 31.7 | 33.3063 | 29.936 | 32.94 | 31.854 | 27.348 | 30.34 | 32.7 | > 25 |  |  |  |  | 16
| 1 | 4 | 30.1814 | 33.46 | 31.8 | 32.9225 | 30.587 | 33.06 | 32.1457 | 23.152 | 30.59 | 33.3 | > 25 |  |  |  |  | 31.7733
| 1 | 5 | 30.3839 | 32.86 | 32.08 | 32.9442 | 29.845 | 32.84 | 31.8472 | 29.307 | 31.77 | 33.9 | > 25 |  |  |  |  | 29
| 1 | 6 | 30.1911 | 32.36 | 31.69 | 32.7406 | 29.887 | 33.06 | 32.322 | 29.796 | 32.46 | 34.4 | > 25
| Mean |  | 30.3068 | 32.9917 | 31.805 | 33.02 | 30.1973 | 32.9083 | 31.9591 | 28.2343 | 31.4183 | 33.4
| Median |  | 30.3309 | 33.16 | 31.75 | 32.9333 | 30.055 | 32.99 | 31.8506 | 29.412 | 31.6 | 33.2
| Std Dev. |  | 0.10187 | 0.54927 | 0.15834 | 0.36727 | 0.38784 | 0.21339 | 0.22254 | 2.6855 | 0.81357 | 0.63246
| Rel.Std.Dev. |  | 0.3361% | 1.665% | 0.4978% | 1.112% | 1.284% | 0.6484% | 0.6963% | 9.511% | 2.589% | 1.894%
| PDM3 |  | -4.616% | 3.834% | 0.09969% | 3.924% | -4.96% | 3.572% | 0.5848% | -11.14% | -1.117% | 5.12%
| Z-Score (Absolute) |  | 0.95 | 0.67 | 0.05 | 0.68 | 1.01 | 0.62 | 0.05 | 1.29 | 0.28 | 0.91 | NA
| Analytical results for Ga in OREAS 461 (Certified Value 50 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 12 | 14 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 51.8 | 47.4 | 47.1 | 44.3 | 29.3579 | 14.3 | 58 | 44 | [7 empty] | 1
| 1 | 2 | 55.9 | 47.6 | 46.1 | 43.9 | 28.3775 | 14.3 | 60 | 43 | [7 empty] | #N/A
| 1 | 3 | 52.1 | 48 | 47.9 | 45.5 | 28.9499 | 15 | 62 | 60 | [7 empty] | 16
| 1 | 4 | 51.8 | 48.2 | 48.3 | 45.1 | 29.6587 | 16.8 | 62 | 45 | [7 empty] | 50.0389
| 1 | 5 | 56 | 50.2 | 47.5 | 45.7 | 29.7931 | 15.9 | 62 | 46 | [7 empty] | 30
| 1 | 6 | 51.1 | 48.3 | 48.2 | 45 | 29.3417 | 16.5 | 62 | 46
| Mean |  | 53.1167 | 48.2833 | 47.5167 | 44.9167 | 29.2465 | 15.4667 | 61 | 47.3333
| Median |  | 51.95 | 48.1 | 47.7 | 45.05 | 29.3498 | 15.45 | 62 | 45.5
| Std Dev. |  | 2.2194 | 1.0008 | 0.82563 | 0.69402 | 0.5164 | 1.093 | 1.6733 | 6.314
| Rel.Std.Dev. |  | 4.178% | 2.073% | 1.738% | 1.545% | 1.766% | 7.067% | 2.743% | 13.34%
| PDM3 |  | 6.151% | -3.508% | -5.041% | -10.24% | -41.55% | -69.09% | 21.91% | -5.407%
| Z-Score (Absolute) |  | 1.04 | 0.31 | 0.2 | 0.2 | 2.55 | 4.62 | 2.31 | 0.21
| Analytical results for Gd in OREAS 461 (Certified Value 100 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 100.5 | 98.9 | 102 | 91.3 | 91.1 | 103.8 | 93.45 | 108.4167 | 87.22 | 101.5 | 118 | 128 | 114 | 97.1 |  | 1
| 1 | 2 | 100.5 | 107.5 | 97.7 | 89.9 | 91.6 | 99.9 | 94.05 | 106.0644 | 87.35 | 99 | 118 | 122 | 104 | 96.2 |  | 8
| 1 | 3 | 105 | 107 | 96.9 | 91 | 95.4 | 99.8 | 92.86 | 108.1757 | 85.97 | 102 | 120.6 | 128 | 106 | 104 |  | 16
| 1 | 4 | 100.5 | 96.3 | 100 | 93.6 | 95.3 | 103 | 92.41 | 109.052 | 86.06 | 102.5 | 123.2 | 132 | 106 | 96.9 |  | 100.3873
| 1 | 5 | 102.5 | 104 | 100.5 | 93.8 | 95.6 | 101.7 | 93.55 | 108.9717 | 89.29 | 102.5 | 118.9 | 133 | 106 | 99.3 |  | 31
| 1 | 6 | 100.5 | 98.9 | 95.7 | 93.8 | 95.6 | 99.6 | 92.68 | 107.1465 | 84.82 | 104.5 | 118.9 | 130 | 112 | 99.4
| Mean |  | 101.5833 | 102.1 | 98.8 | 92.2333 | 94.1 | 101.3 | 93.1667 | 107.9712 | 86.785 | 102 | 119.6 | 128.8333 | 108 | 98.8167
| Median |  | 100.5 | 101.45 | 98.85 | 92.45 | 95.35 | 100.8 | 93.155 | 108.2962 | 86.64 | 102.25 | 118.9 | 129 | 106 | 98.2
| Std Dev. |  | 1.8552 | 4.7096 | 2.4083 | 1.7096 | 2.1392 | 1.8111 | 0.61821 | 1.1598 | 1.5385 | 1.7889 | 2.003 | 3.92 | 4 | 2.8604
| Rel.Std.Dev. |  | 1.826% | 4.613% | 2.438% | 1.854% | 2.273% | 1.788% | 0.6636% | 1.074% | 1.773% | 1.754% | 1.675% | 3.043% | 3.704% | 2.895%
| PDM3 |  | 1.191% | 1.706% | -1.581% | -8.122% | -6.263% | 0.9092% | -7.193% | 7.555% | -13.55% | 1.607% | 19.14% | 28.34% | 7.583% | -1.565%
| Z-Score (Absolute) |  | 0.02 | 0.1 | 0.23 | 0.87 | 0.69 | 0.02 | 0.78 | 0.67 | 1.4 | 0.09 | 1.74 | 2.72 | 0.68 | 0.22
| Analytical results for Ge in OREAS 461 (Indicative Value 2.5 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 20 | 22 | 23 | [12 empty] | ppm
|  |  | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 1 | < 20 | 3 | [12 empty] | 1
| 1 | 2 | < 1 | < 20 | 2 | [12 empty] | 20
| 1 | 3 | < 1 | < 20 | 3 | [12 empty] | 16
| 1 | 4 | < 1 | < 20 | 2 | [12 empty] | 2.5
| 1 | 5 | < 1 | < 20 | 2 | [12 empty] | 21
| 1 | 6 | < 1 | < 20 | 3
| Mean |  |  |  | 2.5
| Median |  |  |  | 2.5
| Std Dev. |  |  |  | 0.54772
| Rel.Std.Dev. |  |  |  | 21.91%
| PDM3 |  |  |  | 0%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Hf in OREAS 461 (Certified Value 14.1 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 20 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 13.5 | 14.9 | 13.1 | 14.2 | 14 | 14.4 | 11.6 | 11.41 | 13 | 15 | 14 |  |  |  |  | 1
| 1 | 2 | 13.9 | 17 | 13.2 | 13.8 | 14.1 | 14.4 | 11.4 | 11.73 | 14 | 14 | 13 |  |  |  |  | 9
| 1 | 3 | 14.1 | 15.1 | 12.9 | 13.8 | 13.8 | 14.4 | 11.4 | 11.01 | 14 | 14 | 15 |  |  |  |  | 16
| 1 | 4 | 13.9 | 14.6 | 13.8 | 14.6 | 14.5 | 14.3 | 11.8 | 11.14 | 14 | 15 | 14 |  |  |  |  | 14.093
| 1 | 5 | 13.4 | 15.5 | 13.9 | 13.8 | 14.1 | 14.2 | 11.4 | 11.08 | 14 | 15 | 14 |  |  |  |  | 32
| 1 | 6 | 13.7 | 15 | 13.1 | 14.3 | 13.8 | 13.9 | 11.7 | 10.75 | 14 | 15 | 13
| Mean |  | 13.75 | 15.35 | 13.3333 | 14.0833 | 14.05 | 14.2667 | 11.55 | 11.1867 | 13.8333 | 14.6667 | 13.8333
| Median |  | 13.8 | 15.05 | 13.15 | 14 | 14.05 | 14.35 | 11.5 | 11.11 | 14 | 15 | 14
| Std Dev. |  | 0.26646 | 0.85965 | 0.41312 | 0.33714 | 0.25884 | 0.19664 | 0.17607 | 0.34074 | 0.40825 | 0.5164 | 0.75277
| Rel.Std.Dev. |  | 1.938% | 5.6% | 3.098% | 2.394% | 1.842% | 1.378% | 1.524% | 3.046% | 2.951% | 3.521% | 5.442%
| PDM3 |  | -2.434% | 8.92% | -5.39% | -0.06833% | -0.3049% | 1.233% | -18.04% | -20.62% | -1.842% | 4.071% | -1.842%
| Z-Score (Absolute) |  | 0.13 | 1.85 | 0.78 | 0.39 | 0.34 | 0.67 | 3.55 | 4.12 | 0 | 1.3 | 0
| Analytical results for Ho in OREAS 461 (Certified Value 4.56 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.2 | 4.49 | 4.52 | 4.54 | 4.29 | 4.4602 | 4.8 | 4.5 | 2.799 | 3.69 | 4.4 | 4.3715 | 4.98 | 5.6 | 4.23 | 1
| 1 | 2 | 4.09 | 4.8 | 4.72 | 4.39 | 4.38 | 4.7545 | 4.8 | 4.5 | 3.2879 | 3.62 | 4.6 | 4.2541 | 4.78 | 4.8 | 4.14 | 10
| 1 | 3 | 4.37 | 4.69 | 4.62 | 4.59 | 4.4 | 4.7341 | 4.9 | 4.4 | 2.9688 | 3.55 | 4.5 | 4.1609 | 4.89 | 5 | 4.86 | 16
| 1 | 4 | 4.24 | 4.46 | 4.76 | 4.65 | 4.36 | 4.8512 | 4.7 | 4.5 | 3.012 | 3.53 | 4.3 | 4.2379 | 5.16 | 5.2 | 4.19 | 4.5595
| 1 | 5 | 4.25 | 4.86 | 4.65 | 4.63 | 4.42 | 4.7886 | 4.8 | 4.4 | 3.1634 | 3.66 | 4.5 | 3.9865 | 5.14 | 5.2 | 3.91 | 33
| 1 | 6 | 4.12 | 4.46 | 4.61 | 4.69 | 4.45 | 4.5705 | 4.6 | 4.4 | 3.1356 | 3.58 | 4.7 | 4.1652 | 5 | 5.4 | 4.06
| Mean |  | 4.2117 | 4.6267 | 4.6467 | 4.5817 | 4.3833 | 4.6932 | 4.7667 | 4.45 | 3.0611 | 3.605 | 4.5 | 4.196 | 4.9917 | 5.2 | 4.2317
| Median |  | 4.22 | 4.59 | 4.635 | 4.61 | 4.39 | 4.7443 | 4.8 | 4.45 | 3.0738 | 3.6 | 4.5 | 4.2016 | 4.99 | 5.2 | 4.165
| Std Dev. |  | 0.10068 | 0.18041 | 0.085245 | 0.10704 | 0.055377 | 0.14755 | 0.10328 | 0.054772 | 0.17151 | 0.062849 | 0.14142 | 0.12804 | 0.14538 | 0.28284 | 0.32799
| Rel.Std.Dev. |  | 2.391% | 3.899% | 1.835% | 2.336% | 1.263% | 3.144% | 2.167% | 1.231% | 5.603% | 1.743% | 3.143% | 3.052% | 2.913% | 5.439% | 7.751%
| PDM3 |  | -7.629% | 1.473% | 1.912% | 0.4861% | -3.864% | 2.932% | 4.544% | -2.402% | -32.86% | -20.93% | -1.305% | -7.972% | 9.478% | 14.05% | -7.19%
| Z-Score (Absolute) |  | 0.73 | 0.32 | 0.37 | 0.21 | 0.3 | 0.49 | 0.67 | 0.13 | 3.64 | 2.26 | 0 | 0.77 | 1.24 | 1.57 | 1
| Analytical results for In in OREAS 461 (Certified Value 0.6 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.63352 | 0.6 | 0.5 | 0.4 | 0.6 | 0.6 | [9 empty] | 1
| 1 | 2 | 0.501 | 0.6 | 0.5 | 0.3 | 0.6 | 0.7 | [9 empty] | #N/A
| 1 | 3 | 0.61738 | 0.6 | 0.5 | 0.3 | 0.6 | 0.6 | [9 empty] | 16
| 1 | 4 | 0.64719 | 0.6 | 0.5 | 0.5 | 0.6 | 0.6 | [9 empty] | 0.60113
| 1 | 5 | 0.69894 | 0.7 | 0.5 | 0.4 | 0.6 | 0.7 | [9 empty] | 34
| 1 | 6 | 0.53581 | 0.7 | 0.5 | 0.3 | 0.6 | 0.8
| Mean |  | 0.60564 | 0.63333 | 0.5 | 0.36667 | 0.6 | 0.66667
| Median |  | 0.62545 | 0.6 | 0.5 | 0.35 | 0.6 | 0.65
| Std Dev. |  | 0.073721 | 0.05164 | 0 | 0.08165 | 0 | 0.08165
| Rel.Std.Dev. |  | 12.17% | 8.154% | 0% | 22.27% | 0% | 12.25%
| PDM3 |  | 0.7505% | 5.358% | -16.82% | -39% | -0.1876% | 10.9%
| Z-Score (Absolute) |  | 0.04 | 0.44 | 1.47 | 3.38 | 0.04 | 0.91
| Analytical results for K in OREAS 461 (Certified Value 0.237 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.298 | 0.27 | 0.23 | 0.28365 | 0.232 | 0.2 | 0.19824 | 0.249 | 0.24 | 0.2 | [5 empty] | 1
| 1 | 2 | 0.281 | 0.25 | 0.23 | 0.27419 | 0.232 | 0.2 | 0.19907 | 0.249 | 0.23 | 0.2 | [5 empty] | 11
| 1 | 3 | 0.281 | 0.26 | 0.22 | 0.28912 | 0.232 | 0.2 | 0.19841 | 0.249 | 0.2 | 0.2 | [5 empty] | 16
| 1 | 4 | 0.295 | 0.28 | 0.23 | 0.2868 | 0.241 | 0.2 | 0.20115 | 0.249 | 0.2 | 0.2 | [5 empty] | 0.23655
| 1 | 5 | 0.281 | 0.27 | 0.23 | 0.2802 | 0.224 | 0.2 | 0.195 | 0.232 | 0.23 | 0.2 | [5 empty] | 35
| 1 | 6 | 0.285 | 0.27 | 0.23 | 0.28052 | 0.224 | 0.2 | 0.19509 | 0.249 | 0.24 | 0.2
| Mean |  | 0.28683 | 0.26667 | 0.22833 | 0.28241 | 0.23083 | 0.2 | 0.19783 | 0.24617 | 0.22333 | 0.2
| Median |  | 0.283 | 0.27 | 0.23 | 0.28208 | 0.232 | 0.2 | 0.19833 | 0.249 | 0.23 | 0.2
| Std Dev. |  | 0.007705 | 0.010328 | 0.0040825 | 0.005328 | 0.0063377 | 0 | 0.0023909 | 0.0069402 | 0.018619 | 0
| Rel.Std.Dev. |  | 2.686% | 3.873% | 1.788% | 1.887% | 2.746% | 0% | 1.209% | 2.819% | 8.337% | 0%
| PDM3 |  | 21.26% | 12.73% | -3.474% | 19.39% | -2.417% | -15.45% | -16.37% | 4.065% | -5.588% | -15.45%
| Z-Score (Absolute) |  | 1.3 | 0.84 | 0.03 | 1.2 | 0.03 | 0.67 | 0.72 | 0.44 | 0.14 | 0.67
| Analytical results for La in OREAS 461 (Certified Value 2690 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 2700 | 2540 | 2710 | 2620 | 2420 | 2420.935 | 2902.1 | 2657 | 2494.0875 | 2081 | 2627 | 2935.0156 | 2670 | 2770 | 3010 | 1
| 1 | 2 | 2690 | 2730 | 2650 | 2540 | 2470 | 2437.2693 | 2797.3 | 2667 | 2471.5684 | 2070 | 2663 | 2981.5177 | 2560 | 2740 | 2970 | #N/A
| 1 | 3 | 2770 | 2610 | 2700 | 2620 | 2550 | 2457.9978 | 2850.5 | 2651 | 2469.9044 | 2025 | 2695 | 2804.781 | 2630 | 2760 | 3060 | 16
| 1 | 4 | 2710 | 2560 | 2760 | 2620 | 2540 | 2425.6309 | 2804.9 | 2662 | 2576.263 | 2009 | 2676 | 2908.064 | 2720 | 2750 | 3000 | 2690.0867
| 1 | 5 | 2690 | 2750 | 2740 | 2640 | 2490 | 2484.4265 | 2809.6 | 2667 | 2558.2758 | 2035 | 2679 | 2853.5237 | 2730 | 2810 | 3100 | 36
| 1 | 6 | 2680 | 2610 | 2780 | 2660 | 2500 | 2472.8625 | 2768 | 2651 | 2568.6476 | 1994 | 2777 | 2893.1142 | 2700 | 2880 | 2970
| Mean |  | 2706.6667 | 2633.3333 | 2723.3333 | 2616.6667 | 2495 | 2449.8537 | 2822.0667 | 2659.1667 | 2523.1245 | 2035.6667 | 2686.1667 | 2896.0027 | 2668.3333 | 2785 | 3018.3333
| Median |  | 2695 | 2610 | 2725 | 2620 | 2495 | 2447.6336 | 2807.25 | 2659.5 | 2526.1816 | 2030 | 2677.5 | 2900.5891 | 2685 | 2765 | 3005
| Std Dev. |  | 32.6599 | 87.3308 | 46.7618 | 40.8248 | 47.6445 | 25.9708 | 47.3297 | 7.3326 | 49.9304 | 34.0451 | 50.0416 | 61.8296 | 64.3169 | 52.4404 | 51.9294
| Rel.Std.Dev. |  | 1.207% | 3.316% | 1.717% | 1.56% | 1.91% | 1.06% | 1.677% | 0.2757% | 1.979% | 1.672% | 1.863% | 2.135% | 2.41% | 1.883% | 1.72%
| PDM3 |  | 0.6163% | -2.11% | 1.236% | -2.729% | -7.252% | -8.93% | 4.906% | -1.149% | -6.207% | -24.33% | -0.1457% | 7.655% | -0.8087% | 3.528% | 12.2%
| Z-Score (Absolute) |  | 0.27 | 0.24 | 0.38 | 0.25 | 1.19 | 1.5 | 1.06 | 0.06 | 1 | 4.35 | 0 | 1.57 | 0 | 0.67 | 2.41
| Analytical results for Li in OREAS 461 (Certified Value 12 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 1
1
| 1 | 1 | 13.2694 | 13 | 12 | 10 | 12 | 10 | [9 empty] | 1
| 1 | 2 | 13.1907 | 13 | 13 | 10 | 13 | 10 | [9 empty] | #N/A
| 1 | 3 | 13.7971 | 13 | 12 | 10 | 12 | 10 | [9 empty] | 16
| 1 | 4 | 12.856 | 13 | 14 | < 10 | 13 | 10 | [9 empty] | 12.0055
| 1 | 5 | 14.3376 | 13 | 13 | 10 | 13 | 10 | [9 empty] | 37
| 1 | 6 | 14.7479 | 13 | 12 | 10 | 13 | 10
| Mean |  | 13.6998 | 13 | 12.6667 | 10 | 12.6667 | 10
| Median |  | 13.5332 | 13 | 12.5 | 10 | 13 | 10
| Std Dev. |  | 0.73091 | 0 | 0.8165 | 0 | 0.5164 | 0
| Rel.Std.Dev. |  | 5.335% | 0% | 6.446% | 0% | 4.077% | 0%
| PDM3 |  | 14.11% | 8.284% | 5.507% | -16.7% | 5.507% | -16.7%
| Z-Score (Absolute) |  | 1.02 | 0.33 | 0 | 2.63 | 0 | 2.63
| Analytical results for Lu in OREAS 461 (Certified Value 0.52 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.52 | 0.5 | 0.49 | 0.52 | 0.49 | 0.5192 | 0.5 | 0.5 | 0.39 | 0.5 | < 1 | 0.3 | 0.6 | 0.69 |  | 1
| 1 | 2 | 0.52 | 0.63 | 0.5 | 0.51 | 0.48 | 0.49987 | 0.5 | 0.5 | 0.4 | 0.5 | < 1 | 0.29 | 0.6 | 0.69 |  | 1
| 1 | 3 | 0.56 | 0.55 | 0.45 | 0.52 | 0.46 | 0.55745 | 0.5 | 0.5 | 0.39 | 0.5 | < 1 | 0.29 | 0.6 | 0.63 |  | 16
| 1 | 4 | 0.54 | 0.54 | 0.5 | 0.55 | 0.49 | 0.5052 | 0.5 | 0.5 | 0.4 | 0.5 | < 1 | 0.42 | 0.6 | 0.83 |  | 0.5154
| 1 | 5 | 0.51 | 0.55 | 0.48 | 0.52 | 0.48 | 0.4725 | 0.6 | 0.4 | 0.4 | 0.5 | < 1 | 0.31 | 0.6 | 0.5 |  | 38
| 1 | 6 | 0.53 | 0.56 | 0.48 | 0.55 | 0.5 | 0.47969 | 0.4 | 0.5 | 0.38 | 0.5 | < 1 | 0.3 | 0.6 | 0.44
| Mean |  | 0.53 | 0.555 | 0.48333 | 0.52833 | 0.48333 | 0.50565 | 0.5 | 0.48333 | 0.39333 | 0.5 |  | 0.31833 | 0.6 | 0.63
| Median |  | 0.525 | 0.55 | 0.485 | 0.52 | 0.485 | 0.50253 | 0.5 | 0.5 | 0.395 | 0.5 |  | 0.3 | 0.6 | 0.66
| Std Dev. |  | 0.017889 | 0.042308 | 0.018619 | 0.017224 | 0.013663 | 0.030571 | 0.063246 | 0.040825 | 0.008165 | 0 |  | 0.050365 | 0 | 0.14156
| Rel.Std.Dev. |  | 3.375% | 7.623% | 3.852% | 3.26% | 2.827% | 6.046% | 12.65% | 8.447% | 2.076% | 0% |  | 15.82% | 0% | 22.47%
| PDM3 |  | 2.833% | 7.684% | -6.221% | 2.51% | -6.221% | -1.891% | -2.988% | -6.221% | -23.68% | -2.988% |  | -38.24% | 16.41% | 22.24%
| Z-Score (Absolute) |  | 0.71 | 0.95 | 0.4 | 0.67 | 0.4 | 0.13 | 0 | 0.4 | 2.54 | 0 | NA | 4.81 | 2.38 | 3.09
| Analytical results for Mg in OREAS 461 (Certified Value 1.05 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.0502 | 1.14 | 1.07 | 1.0292 | 0.947 | 1.105 | 0.99566 | 1.013 | 1.08 | 1.09 | 1.07 |  |  |  |  | 1
| 1 | 2 | 1.0564 | 1.12 | 1.09 | 1.0414 | 0.971 | 1.09 | 0.9842 | 1.007 | 1.05 | 1.09 | 1.06 |  |  |  |  | 2
| 1 | 3 | 1.0553 | 1.12 | 1.05 | 1.035 | 0.953 | 1.09 | 0.99626 | 1.019 | 1 | 1.07 | 1.02 |  |  |  |  | 16
| 1 | 4 | 1.0295 | 1.13 | 1.09 | 1.0208 | 0.965 | 1.095 | 0.98782 | 0.995 | 1 | 1.09 | 1.08 |  |  |  |  | 1.047
| 1 | 5 | 1.0429 | 1.12 | 1.08 | 1.0231 | 0.947 | 1.085 | 0.97274 | 0.977 | 1.03 | 1.11 | 1.09 |  |  |  |  | 39
| 1 | 6 | 1.0377 | 1.1 | 1.05 | 1.0317 | 0.935 | 1.095 | 0.99445 | 1.013 | 1.07 | 1.12 | 1.08
| Mean |  | 1.0453 | 1.1217 | 1.0717 | 1.0302 | 0.953 | 1.0933 | 0.98852 | 1.004 | 1.0383 | 1.095 | 1.0667
| Median |  | 1.0465 | 1.12 | 1.075 | 1.0304 | 0.95 | 1.0925 | 0.99113 | 1.01 | 1.04 | 1.09 | 1.075
| Std Dev. |  | 0.010575 | 0.013292 | 0.018348 | 0.0076445 | 0.013145 | 0.0068313 | 0.0091018 | 0.015531 | 0.034303 | 0.017607 | 0.025033
| Rel.Std.Dev. |  | 1.012% | 1.185% | 1.712% | 0.7421% | 1.379% | 0.6248% | 0.9207% | 1.547% | 3.304% | 1.608% | 2.347%
| PDM3 |  | -0.1595% | 7.132% | 2.356% | -1.606% | -8.978% | 4.426% | -5.585% | -4.107% | -0.8275% | 4.585% | 1.879%
| Z-Score (Absolute) |  | 0 | 1.25 | 0.43 | 0.25 | 1.51 | 0.79 | 0.93 | 0.67 | 0.11 | 0.81 | 0.5
| Analytical results for Mn in OREAS 461 (Certified Value 653 ppm) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [5 empty] | ppm
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | < 2000 | 700 | 695.052 | 540 | 700 | 628.9 | 620 | 680 | 660 | 630 | [5 empty] | 1
| 1 | 2 | < 2000 | 700 | 709.556 | 620 | 700 | 616.5 | 620 | 670 | 660 | 630 | [5 empty] | 3
| 1 | 3 | < 2000 | 700 | 705.802 | 540 | 700 | 625.8 | 540 | 640 | 660 | 660 | [5 empty] | 16
| 1 | 4 | < 2000 | 700 | 700.873 | 620 | 700 | 621.1 | 460 | 640 | 680 | 650 | [5 empty] | 652.6366
| 1 | 5 | < 2000 | 700 | 698.767 | 540 | 700 | 608.7 | 620 | 670 | 680 | 640 | [5 empty] | 40
| 1 | 6 | < 2000 | 700 | 700.313 | 540 | 700 | 618.8 | 620 | 690 | 700 | 640
| Mean |  |  | 700 | 701.7272 | 566.6667 | 700 | 619.9667 | 580 | 665 | 673.3333 | 641.6667
| Median |  |  | 700 | 700.593 | 540 | 700 | 619.95 | 620 | 670 | 670 | 640
| Std Dev. |  |  | 1.2454e-13 | 5.1758 | 41.3118 | 0 | 7.1475 | 66.9328 | 20.7364 | 16.3299 | 11.6905
| Rel.Std.Dev. |  |  | 1.779e-14% | 0.7376% | 7.29% | 0% | 1.153% | 11.54% | 3.118% | 2.425% | 1.822%
| PDM3 |  |  | 7.257% | 7.522% | -13.17% | 7.257% | -5.006% | -11.13% | 1.894% | 3.171% | -1.681%
| Z-Score (Absolute) |  | NA | 0.67 | 0.71 | 1.87 | 0.67 | 0.87 | 1.17 | 0 | 0.16 | 0.45
| Analytical results for Mo in OREAS 461 (Certified Value 45.7 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 1
1
| 1 | 1 | 47 | 24.8881 | 37 | 50 | 39.1087 | 46 | 50 | 48 | [7 empty] | 1
| 1 | 2 | 49 | 24.6069 | 37 | 50 | 41.6616 | 45 | 50 | 45 | [7 empty] | #N/A
| 1 | 3 | 47 | 23.9812 | 37 | 50 | 39.286 | 46 | 50 | 46 | [7 empty] | 16
| 1 | 4 | 47 | 24.9962 | 37 | 50 | 39.9942 | 46 | 50 | 49 | [7 empty] | 45.6795
| 1 | 5 | 48 | 24.4595 | 37 | 50 | 41.1223 | 47 | 50 | 54 | [7 empty] | 41
| 1 | 6 | 47 | 24.441 | 37 | 50 | 43.3661 | 46 | 50 | 49
| Mean |  | 47.5 | 24.5622 | 37 | 50 | 40.7565 | 46 | 50 | 48.5
| Median |  | 47 | 24.5332 | 37 | 50 | 40.5583 | 46 | 50 | 48.5
| Std Dev. |  | 0.83666 | 0.36313 | 0 | 0 | 1.6262 | 0.63246 | 0 | 3.1464
| Rel.Std.Dev. |  | 1.761% | 1.478% | 0% | 0% | 3.99% | 1.375% | 0% | 6.487%
| PDM3 |  | 3.985% | -46.23% | -19% | 9.458% | -10.78% | 0.7016% | 9.458% | 6.175%
| Z-Score (Absolute) |  | 0.16 | 4.6 | 2.02 | 0.67 | 1.24 | 0.16 | 0.67 | 0.36
| Analytical results for Na in OREAS 461 (Indicative Value 0.166 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 11 | 14 | 18 | 20 | [11 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES | BF*OES | [11 empty] | 3
3
| 1 | 1 | 0.17 | 0.171 | 0.15898 | 0.86 | [11 empty] | 1
| 1 | 2 | 0.17 | 0.178 | 0.16647 | 0.85 | [11 empty] | 4
| 1 | 3 | 0.17 | 0.171 | 0.15215 | 0.86 | [11 empty] | 16
| 1 | 4 | 0.17 | 0.178 | 0.15386 | 0.84 | [11 empty] | 0.16624
| 1 | 5 | 0.17 | 0.171 | 0.14555 | 0.77 | [11 empty] | 22
| 1 | 6 | 0.17 | 0.171 | 0.15675 | 0.83
| Mean |  | 0.17 | 0.17333 | 0.15563 | 0.835
| Median |  | 0.17 | 0.171 | 0.1553 | 0.845
| Std Dev. |  | 0 | 0.0036148 | 0.0070257 | 0.033912
| Rel.Std.Dev. |  | 0% | 2.085% | 4.515% | 4.061%
| PDM3 |  | 2.261% | 4.267% | -6.385% | 402.3%
| Z-Score (Absolute) |  | 0.12 | 0.12 | 1.23 | 52.22
| Analytical results for Nb in OREAS 461 (Certified Value 1296 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 1355 | 1365 | 1290 | 1355 | 1215 | 1370 | 1289 | > 1000 | 1110 | 1152.5495 | 1060 | 1310 | 1380 |  |  | 1
| 1 | 2 | 1350 | 1480 | 1300 | 1305 | 1230 | 1375 | 1302 | > 1000 | 1140 | 1160.7615 | 999 | 1310 | 1350 |  |  | 5
| 1 | 3 | 1365 | 1400 | 1280 | 1350 | 1275 | 1406 | 1282 | > 1000 | 1110 | 1039.3668 | 964 | 1310 | 1450 |  |  | 16
| 1 | 4 | 1335 | 1355 | 1330 | 1355 | 1270 | 1387 | 1292 | > 1000 | 1140 | 1146.2002 | 876 | 1260 | 1380 |  |  | 1296.3889
| 1 | 5 | 1335 | 1440 | 1320 | 1350 | 1245 | 1369 | 1296 | > 1000 | 1205 | 1105.055 | 1070 | 1320 | 1390 |  |  | 42
| 1 | 6 | 1350 | 1345 | 1340 | 1370 | 1235 | 1370 | 1271 | -- | 1060 | 1104.737 | 1050 | 1360 | 1390
| Mean |  | 1348.3333 | 1397.5 | 1310 | 1347.5 | 1245 | 1379.5 | 1288.6667 |  | 1127.5 | 1118.1117 | 1003.1667 | 1311.6667 | 1390
| Median |  | 1350 | 1382.5 | 1310 | 1352.5 | 1240 | 1372.5 | 1290.5 |  | 1125 | 1125.6276 | 1024.5 | 1310 | 1385
| Std Dev. |  | 11.6905 | 53.362 | 23.6643 | 22.0794 | 23.4521 | 14.6253 | 10.9484 |  | 47.9322 | 45.4842 | 74.314 | 31.8852 | 32.8634
| Rel.Std.Dev. |  | 0.867% | 3.818% | 1.806% | 1.639% | 1.884% | 1.06% | 0.8496% |  | 4.251% | 4.068% | 7.408% | 2.431% | 2.364%
| PDM3 |  | 4.007% | 7.799% | 1.05% | 3.943% | -3.964% | 6.411% | -0.5957% |  | -13.03% | -13.75% | -22.62% | 1.178% | 7.221%
| Z-Score (Absolute) |  | 0.38 | 0.88 | 0.01 | 0.46 | 0.67 | 0.7 | 0.22 | NA | 1.86 | 1.95 | 3.12 | 0.01 | 0.68
| Analytical results for Nd in OREAS 461 (Certified Value 1629 ppm) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1450 | 1620 | 1665 | 1655 | 1505 | 1612.653 | 1726 | 1610 | 1360.7819 | 1369.17 | 1625 | 1549.4298 | 1650 | 1690 | 1820 | 1
| 1 | 2 | 1460 | 1735 | 1645 | 1600 | 1550 | 1606.5968 | 1708 | 1589.7 | 1412.6371 | 1364.88 | 1648 | 1556.4324 | 1580 | 1520 | 1820 | #N/A
| 1 | 3 | 1525 | 1625 | 1675 | 1640 | 1590 | 1622.9715 | 1706 | 1583.1 | 1363.5954 | 1340.02 | 1653 | 1525.712 | 1640 | 1560 | 1970 | 16
| 1 | 4 | 1485 | 1625 | 1660 | 1645 | 1580 | 1607.0014 | 1698 | 1626.6 | 1374.6739 | 1328.02 | 1654 | 1549.2202 | 1690 | 1550 | 1810 | 1628.7803
| 1 | 5 | 1450 | 1735 | 1680 | 1650 | 1595 | 1658.6709 | 1709 | 1591.1 | 1401.4251 | 1362.31 | 1660 | 1536.8441 | 1690 | 1590 | 1830 | 43
| 1 | 6 | 1440 | 1660 | 1640 | 1670 | 1600 | 1639.0217 | 1669 | 1567.4 | 1342.659 | 1328.88 | 1693 | 1525.046 | 1680 | 1630 | 1850
| Mean |  | 1468.3333 | 1666.6667 | 1660.8333 | 1643.3333 | 1570 | 1624.4859 | 1702.6667 | 1594.65 | 1375.9621 | 1348.88 | 1655.5 | 1540.4474 | 1655 | 1590 | 1850
| Median |  | 1455 | 1642.5 | 1662.5 | 1647.5 | 1585 | 1617.8123 | 1707 | 1590.4 | 1369.1347 | 1351.165 | 1653.5 | 1543.0321 | 1665 | 1575 | 1825
| Std Dev. |  | 31.728 | 54.8331 | 15.9426 | 23.5938 | 36.4692 | 20.7265 | 18.8644 | 20.8302 | 26.4104 | 18.7707 | 22.0068 | 13.2701 | 42.3084 | 61.6441 | 60.3324
| Rel.Std.Dev. |  | 2.161% | 3.29% | 0.9599% | 1.436% | 2.323% | 1.276% | 1.108% | 1.306% | 1.919% | 1.392% | 1.329% | 0.8614% | 2.556% | 3.877% | 3.261%
| PDM3 |  | -9.851% | 2.326% | 1.968% | 0.8935% | -3.609% | -0.2637% | 4.536% | -2.095% | -15.52% | -17.18% | 1.64% | -5.423% | 1.61% | -2.381% | 13.58%
| Z-Score (Absolute) |  | 2.72 | 0.69 | 0.59 | 0.31 | 0.67 | 0 | 1.27 | 0.48 | 4.04 | 4.48 | 0.5 | 1.37 | 0.5 | 0.56 | 3.28
| Analytical results for Ni in OREAS 461 (Certified Value 64 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
1
| 1 | 1 | 69 | 63 | < 100 | 67.7206 | 62 | 70 | 100 | 62 | [7 empty] | 1
| 1 | 2 | 67 | 65 | < 100 | 72.0199 | 60 | 60 | 100 | 66 | [7 empty] | #N/A
| 1 | 3 | 63 | 62 | < 100 | 63.5223 | 56 | 60 | 100 | 74 | [7 empty] | 16
| 1 | 4 | 62 | 69 | < 100 | 63.434 | 52 | 60 | 100 | 70 | [7 empty] | 63.6945
| 1 | 5 | 72 | 62 | < 100 | 60.9882 | 62 | 60 | 100 | 61 | [7 empty] | 44
| 1 | 6 | 48 | 64 | < 100 | 70.3176 | 62 | 70 | 100 | 62
| Mean |  | 63.5 | 64.1667 |  | 66.3338 | 59 | 63.3333 | 100 | 65.8333
| Median |  | 65 | 63.5 |  | 65.6214 | 61 | 60 | 100 | 64
| Std Dev. |  | 8.4558 | 2.6394 |  | 4.3592 | 4.1473 | 5.164 | 1.5567e-14 | 5.2313
| Rel.Std.Dev. |  | 13.32% | 4.113% |  | 6.572% | 7.029% | 8.154% | 1.557e-14% | 7.946%
| PDM3 |  | -0.3054% | 0.7413% |  | 4.144% | -7.37% | -0.5671% | 57% | 3.358%
| Z-Score (Absolute) |  | 0.27 | 0 | NA | 0.88 | 2.09 | 0.34 | 14.5 | 0.67
| Analytical results for P in OREAS 461 (Certified Value 0.359 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.382 | 0.38 | 0.37 | 0.38542 | 0.262 | 0.42 | 0.36324 | 0.33 | 0.34 | 0.33 | [5 empty] | 1
| 1 | 2 | 0.381 | 0.38 | 0.37 | 0.38033 | 0.266 | 0.345 | 0.36442 | 0.32 | 0.34 | 0.33 | [5 empty] | 6
| 1 | 3 | 0.372 | 0.37 | 0.38 | 0.38119 | 0.262 | 0.355 | 0.36721 | 0.34 | 0.34 | 0.34 | [5 empty] | 16
| 1 | 4 | 0.363 | 0.38 | 0.36 | 0.38824 | 0.275 | 0.35 | 0.358 | 0.33 | 0.34 | 0.33 | [5 empty] | 0.35872
| 1 | 5 | 0.38 | 0.37 | 0.37 | 0.37005 | 0.271 | 0.345 | 0.37101 | 0.33 | 0.36 | 0.34 | [5 empty] | 45
| 1 | 6 | 0.381 | 0.37 | 0.36 | 0.37189 | 0.249 | 0.365 | 0.3525 | 0.33 | 0.36 | 0.34
| Mean |  | 0.3765 | 0.375 | 0.36833 | 0.37952 | 0.26417 | 0.36333 | 0.36273 | 0.33 | 0.34667 | 0.335
| Median |  | 0.3805 | 0.375 | 0.37 | 0.38076 | 0.264 | 0.3525 | 0.36383 | 0.33 | 0.34 | 0.335
| Std Dev. |  | 0.0075565 | 0.0054772 | 0.0075277 | 0.0072416 | 0.0090203 | 0.028752 | 0.0066147 | 0.0063246 | 0.010328 | 0.0054772
| Rel.Std.Dev. |  | 2.007% | 1.461% | 2.044% | 1.908% | 3.415% | 7.913% | 1.824% | 1.917% | 2.979% | 1.635%
| PDM3 |  | 4.957% | 4.539% | 2.681% | 5.8% | -26.36% | 1.287% | 1.119% | -8.005% | -3.359% | -6.612%
| Z-Score (Absolute) |  | 0.75 | 0.6 | 0.37 | 0.76 | 3.19 | 0.18 | 0.18 | 0.94 | 0.37 | 0.76
| Analytical results for Pb in OREAS 461 (Certified Value 106 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 105 | 104 | 100 | 106.459 | 57 | < 300 | 110 | 100 | [7 empty] | 1
| 1 | 2 | 106 | 100 | 200 | 106.7039 | 51 | < 300 | 110 | 102 | [7 empty] | #N/A
| 1 | 3 | 103 | 100 | 200 | 103.4708 | 51 | < 300 | 120 | 107 | [7 empty] | 16
| 1 | 4 | 108 | 92 | < 100 | 105.2829 | 58 | < 300 | 110 | 111 | [7 empty] | 106.2928
| 1 | 5 | 105 | 99 | 100 | 101.8455 | 51 | < 300 | 120 | 125 | [7 empty] | 46
| 1 | 6 | 104 | 112 | 100 | 104.6215 | 53 | < 300 | 120 | 107
| Mean |  | 105.1667 | 101.1667 | 140 | 104.7306 | 53.5 |  | 115 | 108.6667
| Median |  | 105 | 100 | 100 | 104.9522 | 52 |  | 115 | 107
| Std Dev. |  | 1.7224 | 6.5853 | 54.7723 | 1.8498 | 3.2094 |  | 5.4772 | 8.9144
| Rel.Std.Dev. |  | 1.638% | 6.509% | 39.12% | 1.766% | 5.999% |  | 4.763% | 8.203%
| PDM3 |  | -1.059% | -4.823% | 31.71% | -1.47% | -49.67% |  | 8.192% | 2.233%
| Z-Score (Absolute) |  | 0 | 0.67 | 5.87 | 0.07 | 8.71 | NA | 1.66 | 0.04
| Analytical results for Pr in OREAS 461 (Certified Value 489 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 439 | 488 | 488 | 485 | 459 | 499.1244 | 498 | 512.1 | 362.8098 | 384.54 | 475.5 | 497.7426 | 482 | 560 | 475 | 1
| 1 | 2 | 442 | 523 | 476 | 471 | 480 | 511.8431 | 451 | 516.9 | 365.3663 | 379.29 | 479 | 499.2078 | 459 | 501 | 471 | #N/A
| 1 | 3 | 455 | 495 | 494 | 484 | 492 | 526.0683 | 474 | 508.9 | 360.9114 | 371.21 | 483.5 | 477.0759 | 473 | 511 | 555 | 16
| 1 | 4 | 445 | 493 | 491 | 490 | 485 | 511.9883 | 463 | 521.7 | 364.3956 | 368.66 | 485.5 | 486.4061 | 489 | 504 | 469 | 488.5896
| 1 | 5 | 448 | 507 | 498 | 485 | 492 | 529.446 | 510 | 502.5 | 361.2096 | 381.3 | 484.5 | 480.6728 | 487 | 516 | 476 | 47
| 1 | 6 | 443 | 486 | 499 | 494 | 490 | 524.5366 | 476 | 504.1 | 361.4149 | 366.33 | 498 | 486.2789 | 484 | 541 | 490
| Mean |  | 445.3333 | 498.6667 | 491 | 484.8333 | 483 | 517.1678 | 478.6667 | 511.0333 | 362.6846 | 375.2217 | 484.3333 | 487.8973 | 479 | 522.1667 | 489.3333
| Median |  | 444 | 494 | 492.5 | 485 | 487.5 | 518.2625 | 475 | 510.5 | 362.1124 | 375.25 | 484 | 486.3425 | 483 | 513.5 | 475.5
| Std Dev. |  | 5.6095 | 14.0095 | 8.438 | 7.7825 | 12.6491 | 11.5327 | 21.8876 | 7.4131 | 1.8475 | 7.4638 | 7.6855 | 8.933 | 11.2606 | 23.3531 | 33.001
| Rel.Std.Dev. |  | 1.26% | 2.809% | 1.719% | 1.605% | 2.619% | 2.23% | 4.573% | 1.451% | 0.5094% | 1.989% | 1.587% | 1.831% | 2.351% | 4.472% | 6.744%
| PDM3 |  | -8.853% | 2.062% | 0.4933% | -0.7688% | -1.144% | 5.849% | -2.031% | 4.594% | -25.77% | -23.2% | -0.8711% | -0.1417% | -1.963% | 6.872% | 0.1522%
| Z-Score (Absolute) |  | 3.09 | 0.7 | 0.48 | 0 | 0.23 | 2.53 | 0.48 | 2.05 | 9.54 | 8.56 | 0.04 | 0.24 | 0.46 | 2.32 | 0.67
| Analytical results for Rb in OREAS 461 (Certified Value 13.5 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  | 1
2
| 1 | 1 | 13.4 | 13.4 | 13.7 | 12.9 | 12.4 | 13.5921 | 13.7 | 14.5 | 13.8667 | 11.2 | 13 | 11.6 | 19 | 13.2 |  | 1
| 1 | 2 | 12.9 | 14.9 | 14 | 12.6 | 13.2 | 13.6904 | 13.5 | 14.2 | 13.7634 | 10.4 | 14 | 11.2 | 18 | 13 |  | #N/A
| 1 | 3 | 13.7 | 13.6 | 14 | 13.2 | 12.7 | 13.5146 | 13.7 | 14.2 | 13.8084 | 10.7 | 13 | 11.3 | 16.5 | 13.1 |  | 16
| 1 | 4 | 12.5 | 13.4 | 13.9 | 13.1 | 13.4 | 13.456 | 14 | 14.3 | 13.8893 | 10.3 | 14 | 11.9 | 15.5 | 13.5 |  | 13.5154
| 1 | 5 | 12.2 | 13.9 | 14.3 | 13.3 | 13.6 | 13.1719 | 13.1 | 14.2 | 13.6644 | 10.1 | 15 | 11.7 | 15.5 | 14.2 |  | 48
| 1 | 6 | 12.7 | 13.5 | 14 | 13.1 | 12.9 | 13.7205 | 13.7 | 13.9 | 14.0615 | 9.9 | 18 | 11.3 | 16.5 | 13
| Mean |  | 12.9 | 13.7833 | 13.9833 | 13.0333 | 13.0333 | 13.5242 | 13.6167 | 14.2167 | 13.8423 | 10.4333 | 14.5 | 11.5 | 16.8333 | 13.3333
| Median |  | 12.8 | 13.55 | 14 | 13.1 | 13.05 | 13.5534 | 13.7 | 14.2 | 13.8376 | 10.35 | 14 | 11.45 | 16.5 | 13.15
| Std Dev. |  | 0.56214 | 0.57764 | 0.19408 | 0.25033 | 0.45019 | 0.19982 | 0.29944 | 0.19408 | 0.13408 | 0.46332 | 1.8708 | 0.27568 | 1.4024 | 0.46332
| Rel.Std.Dev. |  | 4.358% | 4.191% | 1.388% | 1.921% | 3.454% | 1.477% | 2.199% | 1.365% | 0.9687% | 4.441% | 12.9% | 2.397% | 8.331% | 3.475%
| PDM3 |  | -4.554% | 1.982% | 3.462% | -3.567% | -3.567% | 0.06515% | 0.749% | 5.188% | 2.418% | -22.8% | 7.285% | -14.91% | 24.55% | -1.347%
| Z-Score (Absolute) |  | 0.91 | 0.03 | 0.63 | 0.72 | 0.72 | 0.03 | 0.11 | 0.96 | 0.43 | 4.41 | 0.37 | 2.9 | 4.67 | 0.54
| Analytical results for Re in OREAS 461 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 17 | 22 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | [12 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 1 | [12 empty] | 17
| 1 | 3 | < 0.1 | < 0.1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | [12 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | [12 empty] | 23
| 1 | 6 | < 0.1 | < 0.1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 461 (Certified Value 518 ppm) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 0
0
| 1 | 1 | 410 | 500 | 400 | 700 | 600 | [10 empty] | 1
| 1 | 2 | 450 | 600 | 400 | 600 | 600 | [10 empty] | 10
| 1 | 3 | 420 | 500 | 500 | 600 | 400 | [10 empty] | 16
| 1 | 4 | 410 | 500 | 500 | 600 | 600 | [10 empty] | 517.6667
| 1 | 5 | 460 | 500 | 500 | 500 | 600 | [10 empty] | 49
| 1 | 6 | 460 | < 500 | 500 | 400 | 600
| Mean |  | 435 | 520 | 466.6667 | 566.6667 | 566.6667
| Median |  | 435 | 500 | 500 | 600 | 600
| Std Dev. |  | 24.2899 | 44.7214 | 51.6398 | 103.2796 | 81.6497
| Rel.Std.Dev. |  | 5.584% | 8.6% | 11.07% | 18.23% | 14.41%
| PDM3 |  | -15.97% | 0.4507% | -9.852% | 9.466% | 9.466%
| Z-Score (Absolute) |  | 1.07 | 0 | 0.67 | 0.59 | 1.01
| Analytical results for Sb in OREAS 461 (Certified Value 2.76 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 2.7 | 2.9 | 3 | 2.6073 | 3 | 2.6 | [9 empty] | 1
| 1 | 2 | 2.6 | 2.8 | 3 | 2.5343 | 3 | 2.4 | [9 empty] | #N/A
| 1 | 3 | 2.6 | 2.8 | 3 | 2.3028 | 3 | 2.7 | [9 empty] | 16
| 1 | 4 | 2.7 | 2.7 | 3 | 2.3942 | 3 | 2.8 | [9 empty] | 2.7613
| 1 | 5 | 2.7 | 2.8 | 3 | 2.4103 | 3 | 2.7 | [9 empty] | 50
| 1 | 6 | 2.6 | 2.8 | 3 | 2.3593 | 3 | 2.9
| Mean |  | 2.65 | 2.8 | 3 | 2.4347 | 3 | 2.6833
| Median |  | 2.65 | 2.8 | 3 | 2.4023 | 3 | 2.7
| Std Dev. |  | 0.054772 | 0.063246 | 0 | 0.11405 | 0 | 0.17224
| Rel.Std.Dev. |  | 2.067% | 2.259% | 0% | 4.684% | 0% | 6.419%
| PDM3 |  | -4.032% | 1.4% | 8.643% | -11.83% | 8.643% | -2.825%
| Z-Score (Absolute) |  | 0.35 | 0.22 | 1 | 1.18 | 1 | 0.22
| Analytical results for Sc in OREAS 461 (Indicative Value 44 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | [11 empty] | 1
1
| 1 | 1 | 42 | 50 | 40 | 110 | [11 empty] | 1
| 1 | 2 | 43 | 50 | 40 | 100 | [11 empty] | 18
| 1 | 3 | 42 | 50 | 40 | 120 | [11 empty] | 16
| 1 | 4 | 42 | 50 | 40 | 110 | [11 empty] | 44
| 1 | 5 | 42 | 50 | 40 | 110 | [11 empty] | 24
| 1 | 6 | 41 | 50 | 40 | 110
| Mean |  | 42 | 50 | 40 | 110
| Median |  | 42 | 50 | 40 | 110
| Std Dev. |  | 0.63246 | 0 | 0 | 6.3246
| Rel.Std.Dev. |  | 1.506% | 0% | 0% | 5.75%
| PDM3 |  | -4.545% | 13.64% | -9.091% | 150%
| Z-Score (Absolute) |  | 0.54 | 0.54 | 0.81 | 8.36
| Analytical results for Se in OREAS 461 (Indicative Value < 20 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 1
1
| 1 | 1 | < 20 | < 20 | [13 empty] | 1
| 1 | 2 | < 20 | < 20 | [13 empty] | 15
| 1 | 3 | < 20 | < 20 | [13 empty] | 16
| 1 | 4 | < 20 | < 20 | [13 empty] | < 20
| 1 | 5 | < 20 | < 20 | [13 empty] | 21
| 1 | 6 | < 20 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Si in OREAS 461 (Certified Value 14.71 wt.%) | [16 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 14.9999 | 14.41 | 15.4308 | 14.248 | 14.9 | 14.5004 | 13.323 | 14.4 | [7 empty] | 1
| 1 | 2 | 15.2174 | 14.45 | 15.224 | 14.613 | 14.75 | 14.4504 | 13.369 | 14.4 | [7 empty] | 7
| 1 | 3 | 15.1581 | 14.32 | 14.93 | 14.281 | 14.88 | 14.5304 | 13.136 | 14.1 | [7 empty] | 16
| 1 | 4 | 15.3163 | 14.45 | 15.036 | 14.851 | 14.86 | 14.6505 | 13.136 | 14.4 | [7 empty] | 14.7082
| 1 | 5 | 15.3064 | 14.51 | 15.056 | 14.075 | 14.75 | 14.5104 | 12.949 | 14.7 | [7 empty] | 51
| 1 | 6 | 14.99 | 14.43 | 15.0834 | 14.047 | 14.87 | 14.7555 | 13.182 | 14.9
| Mean |  | 15.1647 | 14.4283 | 15.1267 | 14.3525 | 14.835 | 14.5663 | 13.1825 | 14.4833
| Median |  | 15.1878 | 14.44 | 15.0697 | 14.2645 | 14.865 | 14.5204 | 13.159 | 14.4
| Std Dev. |  | 0.14394 | 0.062743 | 0.17648 | 0.31709 | 0.067157 | 0.11408 | 0.15058 | 0.27869
| Rel.Std.Dev. |  | 0.9492% | 0.4349% | 1.167% | 2.209% | 0.4527% | 0.7832% | 1.142% | 1.924%
| PDM3 |  | 3.104% | -1.902% | 2.846% | -2.418% | 0.8624% | -0.9646% | -10.37% | -1.529%
| Z-Score (Absolute) |  | 1.79 | 0.27 | 1.68 | 0.48 | 0.87 | 0.12 | 3.75 | 0.12
| Analytical results for Sm in OREAS 461 (Certified Value 220 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 214 | 222 | 217 | 218 | 198 | 219.799 | 226.8 | 221 | 172.2022 | 177.6 | 210 | 223.855 | 225 | 240 | 224 | 1
| 1 | 2 | 212 | 238 | 216 | 213 | 200 | 214.5312 | 221.6 | 221.9 | 175.7228 | 176.3 | 211.5 | 227.0854 | 213 | 219 | 218 | 8
| 1 | 3 | 220 | 221 | 221 | 219 | 208 | 215.5925 | 226 | 219.7 | 172.0108 | 173.7 | 212.5 | 216.0076 | 219 | 217 | 248 | 16
| 1 | 4 | 217 | 221 | 220 | 221 | 206 | 216.2806 | 221.1 | 223.6 | 175.0184 | 172 | 213 | 223.3639 | 229 | 217 | 224 | 219.6468
| 1 | 5 | 215 | 238 | 224 | 220 | 209 | 216.4453 | 220.3 | 218.8 | 176.8255 | 175.7 | 217 | 219.3831 | 230 | 226 | 222 | 52
| 1 | 6 | 213 | 222 | 218 | 225 | 209 | 219.4423 | 219 | 219.2 | 171.1857 | 172.9 | 221.5 | 221.9659 | 226 | 233 | 233
| Mean |  | 215.1667 | 227 | 219.3333 | 219.3333 | 205 | 217.0151 | 222.4667 | 220.7 | 173.8276 | 174.7 | 214.25 | 221.9435 | 223.6667 | 225.3333 | 228.1667
| Median |  | 214.5 | 222 | 219 | 219.5 | 207 | 216.3629 | 221.35 | 220.35 | 173.6103 | 174.7 | 212.75 | 222.6649 | 225.5 | 222.5 | 224
| Std Dev. |  | 2.9269 | 8.5323 | 2.9439 | 3.9328 | 4.8166 | 2.1307 | 3.181 | 1.833 | 2.3203 | 2.1679 | 4.2515 | 3.8445 | 6.5013 | 9.5219 | 10.8888
| Rel.Std.Dev. |  | 1.36% | 3.759% | 1.342% | 1.793% | 2.35% | 0.9818% | 1.43% | 0.8306% | 1.335% | 1.241% | 1.984% | 1.732% | 2.907% | 4.226% | 4.772%
| PDM3 |  | -2.04% | 3.348% | -0.1427% | -0.1427% | -6.668% | -1.198% | 1.284% | 0.4795% | -20.86% | -20.46% | -2.457% | 1.046% | 1.83% | 2.589% | 3.879%
| Z-Score (Absolute) |  | 0.65 | 1.19 | 0 | 0 | 2.23 | 0.36 | 0.49 | 0.21 | 7.08 | 6.95 | 0.79 | 0.41 | 0.67 | 0.93 | 0.76
| Analytical results for Sn in OREAS 461 (Certified Value 25.6 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS |  | 1
1
| 1 | 1 | 30 | 26 | 29 | 27 | 26 | 27 | 22 | 21 | 23 | 17.7419 | 103 | < 50 | 40 | 24 |  | 1
| 1 | 2 | 31 | 28 | 28 | 26 | 26 | 29 | 21 | 20.9 | 24 | 17.4891 | 26 | < 50 | 40 | 23 |  | 9
| 1 | 3 | 32 | 26 | 28 | 27 | 27 | 28 | 21 | 19.4 | 25 | 16.8012 | 25 | < 50 | 40 | 28 |  | 16
| 1 | 4 | 30 | 25 | 30 | 27 | 28 | 28 | 21 | 18.7 | 24 | 16.8767 | 24 | < 50 | 30 | 22 |  | 25.6303
| 1 | 5 | 30 | 27 | 28 | 28 | 26 | 28 | 22 | 18.8 | 24 | 18.6538 | 28 | < 50 | 40 | 22 |  | 53
| 1 | 6 | 31 | 26 | 29 | 29 | 27 | 29 | 21 | 18.8 | 25 | 15.9929 | 28 | < 50 | 40 | 23
| Mean |  | 30.6667 | 26.3333 | 28.6667 | 27.3333 | 26.6667 | 28.1667 | 21.3333 | 19.6 | 24.1667 | 17.2592 | 39 |  | 38.3333 | 23.6667
| Median |  | 30.5 | 26 | 28.5 | 27 | 26.5 | 28 | 21 | 19.1 | 24 | 17.1829 | 27 |  | 40 | 23
| Std Dev. |  | 0.8165 | 1.0328 | 0.8165 | 1.0328 | 0.8165 | 0.75277 | 0.5164 | 1.0752 | 0.75277 | 0.91545 | 31.3943 |  | 4.0825 | 2.2509
| Rel.Std.Dev. |  | 2.662% | 3.922% | 2.848% | 3.779% | 3.062% | 2.673% | 2.421% | 5.486% | 3.115% | 5.304% | 80.5% |  | 10.65% | 9.511%
| PDM3 |  | 19.65% | 2.743% | 11.85% | 6.645% | 4.044% | 9.896% | -16.77% | -23.53% | -5.711% | -32.66% | 52.16% |  | 49.56% | -7.661%
| Z-Score (Absolute) |  | 1.25 | 0 | 0.67 | 0.29 | 0.1 | 0.53 | 1.44 | 1.95 | 0.63 | 2.62 | 0.04 | NA | 3.47 | 1.02
| Analytical results for Sr in OREAS 461 (Certified Value 579 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 577 | 558 | 571 | 574 | 566 | 602 | 573.4 | 546 | 600 | 592.9748 | 533 | 583 | 570 |  |  | 1
| 1 | 2 | 574 | 600 | 579 | 554 | 567 | 608 | 571.9 | 571 | 600 | 621.3923 | 503 | 576 | 570 |  |  | 11
| 1 | 3 | 601 | 574 | 590 | 577 | 575 | 603 | 566.9 | 553 | 600 | 604.2155 | 520 | 592 | 540 |  |  | 16
| 1 | 4 | 574 | 560 | 580 | 574 | 570 | 596 | 580.4 | 574 | 600 | 624.9661 | 532 | 581 | 570 |  |  | 578.9394
| 1 | 5 | 575 | 590 | 597 | 578 | 585 | 604 | 569.2 | 540 | 600 | 650.914 | 535 | 593 | 580 |  |  | 54
| 1 | 6 | 575 | 564 | 579 | 578 | 577 | 588 | 567 | 541 | 600 | 673.8734 | 523 | 597 | 580
| Mean |  | 579.3333 | 574.3333 | 582.6667 | 572.5 | 573.3333 | 600.1667 | 571.4667 | 554.1667 | 600 | 628.056 | 524.3333 | 587 | 568.3333
| Median |  | 575 | 569 | 579.5 | 575.5 | 572.5 | 602.5 | 570.55 | 549.5 | 600 | 623.1792 | 527.5 | 587.5 | 570
| Std Dev. |  | 10.6708 | 17.224 | 9.2664 | 9.2466 | 7.174 | 7.111 | 5.0926 | 14.9588 | 0 | 29.9157 | 12.0277 | 8.1731 | 14.7196
| Rel.Std.Dev. |  | 1.842% | 2.999% | 1.59% | 1.615% | 1.251% | 1.185% | 0.8911% | 2.699% | 0% | 4.763% | 2.294% | 1.392% | 2.59%
| PDM3 |  | 0.06805% | -0.7956% | 0.6438% | -1.112% | -0.9683% | 3.667% | -1.291% | -4.279% | 3.638% | 8.484% | -9.432% | 1.392% | -1.832%
| Z-Score (Absolute) |  | 0 | 0.06 | 0.67 | 0.11 | 0.15 | 2.21 | 0.31 | 1.83 | 2.2 | 4.67 | 4.46 | 1.06 | 0.09
| Analytical results for Ta in OREAS 461 (Certified Value 25.1 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 1
2
| 1 | 1 | 23.4 | 23.9 | 23.7 | 24.2 | 25.4 | 24.6551 | 24.5 | 18.43 | 21 | 23.7574 | 28.4 | 25.5 | 25 |  |  | 1
| 1 | 2 | 23.7 | 26.3 | 24.8 | 24.5 | 25.5 | 24.9003 | 24.7 | 18.25 | 22 | 21.205 | 26.7 | 25.5 | 25.9 |  |  | 1
| 1 | 3 | 24.5 | 25.7 | 24.4 | 25.2 | 26.1 | 25.0581 | 25.3 | 17.73 | 17 | 19.8639 | 25.8 | 25 | 30.5 |  |  | 16
| 1 | 4 | 23.9 | 23.1 | 25.4 | 25.1 | 26.3 | 25.1193 | 25.2 | 17.59 | 23 | 21.2055 | 23.8 | 25.5 | 25.2 |  |  | 25.1318
| 1 | 5 | 23.7 | 26.3 | 25.2 | 25.2 | 26.2 | 24.3658 | 25 | 17.92 | 22 | 18.0361 | 28.6 | 26.5 | 25.1 |  |  | 55
| 1 | 6 | 23.6 | 24 | 24.1 | 25.3 | 26 | 24.5507 | 24.5 | 17.68 | 17 | 19.6964 | 28.1 | 27 | 25.8
| Mean |  | 23.8 | 24.8833 | 24.6 | 24.9167 | 25.9167 | 24.7749 | 24.8667 | 17.9333 | 20.3333 | 20.6274 | 26.9 | 25.8333 | 26.25
| Median |  | 23.7 | 24.85 | 24.6 | 25.15 | 26.05 | 24.7777 | 24.85 | 17.825 | 21.5 | 20.5344 | 27.4 | 25.5 | 25.5
| Std Dev. |  | 0.37947 | 1.3862 | 0.65422 | 0.4535 | 0.37639 | 0.29884 | 0.35024 | 0.3378 | 2.6583 | 1.9302 | 1.8676 | 0.75277 | 2.1154
| Rel.Std.Dev. |  | 1.594% | 5.571% | 2.659% | 1.82% | 1.452% | 1.206% | 1.408% | 1.884% | 13.07% | 9.358% | 6.943% | 2.914% | 8.059%
| PDM3 |  | -5.299% | -0.9887% | -2.116% | -0.8561% | 3.123% | -1.42% | -1.055% | -28.64% | -19.09% | -17.92% | 7.036% | 2.791% | 4.449%
| Z-Score (Absolute) |  | 0.98 | 0.02 | 0.25 | 0.05 | 0.97 | 0.08 | 0 | 6.38 | 4.17 | 3.9 | 1.56 | 0.67 | 0.49
| Analytical results for Tb in OREAS 461 (Certified Value 9.08 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 8.63 | 9 | 8.47 | 8.68 | 8.77 | 9.424 | 9.4 | 9.2 | 6.9831 | 8.37 | 9 | 11.7624 | 9.08 | 11 | 9.13 | 1
| 1 | 2 | 8.7 | 9.67 | 9.26 | 8.49 | 8.81 | 9.3774 | 9.4 | 9.1 | 7.203 | 8.29 | 9.3 | 11.8144 | 8.76 | 9.8 | 8.84 | 2
| 1 | 3 | 9.06 | 9.34 | 8.89 | 8.47 | 8.92 | 9.2679 | 9.2 | 9.2 | 7.1559 | 8.11 | 9.3 | 11.4646 | 8.97 | 10 | 9.21 | 16
| 1 | 4 | 8.61 | 8.62 | 8.83 | 9 | 8.99 | 9.2756 | 9.3 | 9.2 | 7.1965 | 8.11 | 9.3 | 11.6219 | 9.62 | 10 | 8.06 | 9.0752
| 1 | 5 | 8.85 | 9.59 | 8.91 | 8.76 | 9.06 | 9.3693 | 9.2 | 8.9 | 7.0072 | 8.28 | 9.2 | 11.4073 | 9.64 | 10.6 | 9.03 | 56
| 1 | 6 | 8.55 | 8.9 | 8.9 | 8.89 | 8.91 | 9.6168 | 9.2 | 9.1 | 7.2321 | 8.06 | 9.5 | 11.5444 | 9.35 | 10.4 | 9.36
| Mean |  | 8.7333 | 9.1867 | 8.8767 | 8.715 | 8.91 | 9.3885 | 9.2833 | 9.1167 | 7.1296 | 8.2033 | 9.2667 | 11.6025 | 9.2367 | 10.3 | 8.9383
| Median |  | 8.665 | 9.17 | 8.895 | 8.72 | 8.915 | 9.3733 | 9.25 | 9.15 | 7.1762 | 8.195 | 9.3 | 11.5832 | 9.215 | 10.2 | 9.08
| Std Dev. |  | 0.19023 | 0.41423 | 0.25153 | 0.21249 | 0.10826 | 0.12741 | 0.098319 | 0.1169 | 0.10727 | 0.1258 | 0.1633 | 0.16206 | 0.35926 | 0.45166 | 0.4643
| Rel.Std.Dev. |  | 2.178% | 4.509% | 2.834% | 2.438% | 1.215% | 1.357% | 1.059% | 1.282% | 1.505% | 1.534% | 1.762% | 1.397% | 3.889% | 4.385% | 5.195%
| PDM3 |  | -3.767% | 1.228% | -2.188% | -3.969% | -1.821% | 3.452% | 2.293% | 0.4566% | -21.44% | -9.607% | 2.109% | 27.85% | 1.779% | 13.5% | -1.508%
| Z-Score (Absolute) |  | 1.08 | 0.2 | 0.67 | 1.13 | 0.58 | 0.76 | 0.47 | 0 | 5.58 | 2.57 | 0.42 | 6.98 | 0.34 | 3.32 | 0.01
| Analytical results for Te in OREAS 461 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 1 | < 1 | [13 empty] | 1
| 1 | 2 | < 1 | < 1 | [13 empty] | 16
| 1 | 3 | < 1 | < 1 | [13 empty] | 16
| 1 | 4 | < 1 | < 1 | [13 empty] | < 1
| 1 | 5 | < 1 | < 1 | [13 empty] | 22
| 1 | 6 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 461 (Certified Value 210 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 201 | 205 | 211 | 204 | 206 | 208.3192 | 221.7 | 210.7 | 157.2708 | 153.7 | 202.5 | 150.0488 | 156 | 221 | 204 | 1
| 1 | 2 | 201 | 220 | 211 | 199 | 215 | 213.6036 | 219.8 | 210 | 156.3772 | 154.7 | 204.5 | 145.9768 | 149 | 220 | 204 | #N/A
| 1 | 3 | 206 | 215 | 214 | 203 | 209 | 210.6586 | 219.8 | 209.5 | 156.9177 | 152.2 | 203.5 | 144.9203 | 150 | 221 | 193 | 16
| 1 | 4 | 202 | 200 | 209 | 207 | 217 | 210.8484 | 219.4 | 211.9 | 155.2582 | 150.6 | 205 | 146.9643 | 157 | 220 | 202 | 209.8552
| 1 | 5 | 203 | 216 | 210 | 207 | 212 | 207.4283 | 221.7 | 206.8 | 153.9432 | 151.9 | 204 | 145.3351 | 157 | 222 | 204 | 57
| 1 | 6 | 200 | 203 | 211 | 209 | 216 | 210.5839 | 216.1 | 206.6 | 155.6776 | 149.7 | 212 | 146.6704 | 156 | 224 | 204
| Mean |  | 202.1667 | 209.8333 | 211 | 204.8333 | 212.5 | 210.2403 | 219.75 | 209.25 | 155.9074 | 152.1333 | 205.25 | 146.6526 | 154.1667 | 221.3333 | 201.8333
| Median |  | 201.5 | 210 | 211 | 205.5 | 213.5 | 210.6212 | 219.8 | 209.75 | 156.0274 | 152.05 | 204.25 | 146.3236 | 156 | 221 | 204
| Std Dev. |  | 2.137 | 8.1833 | 1.6733 | 3.6009 | 4.3243 | 2.1713 | 2.0521 | 2.1342 | 1.2192 | 1.864 | 3.4169 | 1.8344 | 3.656 | 1.5055 | 4.4008
| Rel.Std.Dev. |  | 1.057% | 3.9% | 0.793% | 1.758% | 2.035% | 1.033% | 0.9338% | 1.02% | 0.782% | 1.225% | 1.665% | 1.251% | 2.371% | 0.6802% | 2.18%
| PDM3 |  | -3.664% | -0.01041% | 0.5455% | -2.393% | 1.26% | 0.1835% | 4.715% | -0.2884% | -25.71% | -27.51% | -2.194% | -30.12% | -26.54% | 5.47% | -3.823%
| Z-Score (Absolute) |  | 0.29 | 0.55 | 0.67 | 0 | 0.84 | 0.59 | 1.63 | 0.48 | 5.35 | 5.76 | 0.1 | 6.36 | 5.54 | 1.8 | 0.13
| Analytical results for Ti in OREAS 461 (Certified Value 1.84 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.862 | 1.93 | 1.81 | 1.8676 | 1.733 | 1.88 | 1.7772 | 1.77 | 1.913 | 1.9 | 1.83 |  |  |  |  | 1
| 1 | 2 | 1.878 | 1.9 | 1.81 | 1.8004 | 1.786 | 1.86 | 1.7703 | 1.72 | 1.862 | 1.9 | 1.83 |  |  |  |  | 3
| 1 | 3 | 1.852 | 1.91 | 1.8 | 1.86 | 1.733 | 1.88 | 1.7749 | 1.6 | 1.802 | 1.87 | 1.67 |  |  |  |  | 16
| 1 | 4 | 1.846 | 1.92 | 1.81 | 1.8383 | 1.768 | 1.87 | 1.7869 | 1.37 | 1.827 | 1.88 | 1.84 |  |  |  |  | 1.8399
| 1 | 5 | 1.801 | 1.9 | 1.82 | 1.8314 | 1.715 | 1.85 | 1.7662 | 1.71 | 1.865 | 1.91 | 1.87 |  |  |  |  | 58
| 1 | 6 | 1.871 | 1.87 | 1.8 | 1.8296 | 1.709 | 1.88 | 1.7938 | 1.75 | 1.924 | 1.94 | 1.84
| Mean |  | 1.8517 | 1.905 | 1.8083 | 1.8379 | 1.7407 | 1.87 | 1.7782 | 1.6533 | 1.8655 | 1.9 | 1.8133
| Median |  | 1.857 | 1.905 | 1.81 | 1.8349 | 1.733 | 1.875 | 1.776 | 1.715 | 1.8635 | 1.9 | 1.835
| Std Dev. |  | 0.027471 | 0.020736 | 0.0075277 | 0.02404 | 0.03027 | 0.012649 | 0.010393 | 0.15082 | 0.04734 | 0.024495 | 0.07174
| Rel.Std.Dev. |  | 1.484% | 1.089% | 0.4163% | 1.308% | 1.739% | 0.6764% | 0.5845% | 9.122% | 2.538% | 1.289% | 3.956%
| PDM3 |  | 0.639% | 3.538% | -1.716% | -0.1104% | -5.394% | 1.635% | -3.353% | -10.14% | 1.391% | 3.266% | -1.444%
| Z-Score (Absolute) |  | 0.19 | 1.26 | 0.67 | 0.08 | 2.03 | 0.56 | 1.28 | 2.64 | 0.47 | 1.16 | 0
| Analytical results for Tl in OREAS 461 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 14 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 17
| 1 | 3 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | < 0.5
| 1 | 5 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 23
| 1 | 6 | < 0.5 | < 0.5 | 0.1 | < 0.5 | < 2 | < 0.5
| Mean |  |  |  | 0.1
| Median |  |  |  | 0.1
| Std Dev. |  |  |  | 1.6997e-17
| Rel.Std.Dev. |  |  |  | 1.7e-14%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Tm in OREAS 461 (Certified Value 0.89 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 0.83 | 0.87 | 0.87 | 0.96 | 0.92 | 0.83189 | 1 | 1 | 0.5457 | 0.72 | 0.9 | 0.71262 | 0.9 | 1.4 | 0.82 | 1
| 1 | 2 | 0.84 | 0.92 | 0.86 | 0.88 | 0.89 | 0.80514 | 0.9 | 1 | 0.623 | 0.67 | 0.9 | 0.74202 | 0.84 | 1 | 0.92 | 4
| 1 | 3 | 0.86 | 1.01 | 0.8 | 0.93 | 0.91 | 0.76837 | 0.9 | 1 | 0.58336 | 0.68 | 0.9 | 0.7014 | 0.91 | 1.2 | 0.92 | 16
| 1 | 4 | 0.84 | 0.9 | 0.89 | 0.92 | 0.88 | 0.89232 | 0.9 | 1 | 0.5609 | 0.68 | 0.9 | 0.69532 | 1.12 | 1 | 0.76 | 0.88945
| 1 | 5 | 0.84 | 0.99 | 0.91 | 0.97 | 0.88 | 0.84197 | 1 | 0.9 | 0.62366 | 0.7 | 0.9 | 0.67348 | 0.92 | 1 | 0.71 | 59
| 1 | 6 | 0.84 | 0.87 | 0.82 | 1.01 | 0.9 | 0.89356 | 0.9 | 1 | 0.5576 | 0.67 | 1 | 0.69125 | 0.92 | 1 | 0.8
| Mean |  | 0.84167 | 0.92667 | 0.85833 | 0.945 | 0.89667 | 0.83888 | 0.93333 | 0.98333 | 0.58237 | 0.68667 | 0.91667 | 0.70268 | 0.935 | 1.1 | 0.82167
| Median |  | 0.84 | 0.91 | 0.865 | 0.945 | 0.895 | 0.83693 | 0.9 | 1 | 0.57213 | 0.68 | 0.9 | 0.69836 | 0.915 | 1 | 0.81
| Std Dev. |  | 0.0098319 | 0.060222 | 0.041673 | 0.045056 | 0.01633 | 0.04902 | 0.05164 | 0.040825 | 0.033988 | 0.019664 | 0.040825 | 0.023178 | 0.095446 | 0.16733 | 0.084951
| Rel.Std.Dev. |  | 1.168% | 6.499% | 4.855% | 4.768% | 1.821% | 5.844% | 5.533% | 4.152% | 5.836% | 2.864% | 4.454% | 3.298% | 10.21% | 15.21% | 10.34%
| PDM3 |  | -5.373% | 4.184% | -3.499% | 6.245% | 0.811% | -5.686% | 4.933% | 10.55% | -34.52% | -22.8% | 3.06% | -21% | 5.121% | 23.67% | -7.621%
| Z-Score (Absolute) |  | 0.67 | 0.37 | 0.47 | 0.59 | 0 | 0.71 | 0.45 | 1.06 | 3.85 | 2.57 | 0.25 | 2.38 | 0.02 | 1.27 | 0.92
| Analytical results for U in OREAS 461 (Certified Value 4.79 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.48 | 4.67 | 4.7 | 4.61 | 4.66 | 5.0344 | 4.8 | 4.9 | 3.7646 | 3.6 | 4.5 | 3.9771 | 3.07 | 5 | 5.03 | 1
| 1 | 2 | 4.46 | 5.04 | 4.6 | 4.6 | 4.68 | 4.7932 | 4.8 | 4.9 | 3.8339 | 3.65 | 4.5 | 4.0424 | 2.93 | 4.5 | 4.76 | 5
| 1 | 3 | 4.7 | 4.9 | 4.7 | 4.72 | 4.83 | 5.0458 | 4.8 | 4.9 | 3.8137 | 3.54 | 4.5 | 3.9126 | 3.01 | 4.5 | 5.06 | 16
| 1 | 4 | 4.55 | 4.6 | 4.6 | 4.69 | 4.97 | 4.8958 | 4.8 | 4.8 | 3.6256 | 3.47 | 5.5 | 4.053 | 3.22 | 5 | 5.12 | 4.7882
| 1 | 5 | 4.6 | 4.91 | 4.9 | 4.7 | 4.85 | 4.5699 | 4.8 | 4.9 | 3.8763 | 3.44 | 4.5 | 4.0302 | 3.12 | 5 | 5.26 | 60
| 1 | 6 | 4.52 | 4.6 | 4.9 | 4.83 | 4.94 | 4.8911 | 4.7 | 4.9 | 3.8422 | 3.42 | 5 | 4.1097 | 3.04 | 5 | 5.1
| Mean |  | 4.5517 | 4.7867 | 4.7333 | 4.6917 | 4.8217 | 4.8717 | 4.7833 | 4.8833 | 3.7927 | 3.52 | 4.75 | 4.0208 | 3.065 | 4.8333 | 5.055
| Median |  | 4.535 | 4.785 | 4.7 | 4.695 | 4.84 | 4.8934 | 4.8 | 4.9 | 3.8238 | 3.505 | 4.5 | 4.0363 | 3.055 | 5 | 5.08
| Std Dev. |  | 0.088185 | 0.18737 | 0.13663 | 0.083766 | 0.12891 | 0.17608 | 0.040825 | 0.040825 | 0.089766 | 0.092304 | 0.41833 | 0.068002 | 0.098944 | 0.2582 | 0.16489
| Rel.Std.Dev. |  | 1.937% | 3.914% | 2.886% | 1.785% | 2.673% | 3.614% | 0.8535% | 0.836% | 2.367% | 2.622% | 8.807% | 1.691% | 3.228% | 5.342% | 3.262%
| PDM3 |  | -4.941% | -0.03298% | -1.147% | -2.017% | 0.698% | 1.743% | -0.1026% | 1.986% | -20.79% | -26.49% | -0.7987% | -16.03% | -35.99% | 0.9416% | 5.571%
| Z-Score (Absolute) |  | 0.89 | 0.26 | 0 | 0.2 | 0.43 | 0.67 | 0.24 | 0.73 | 4.42 | 5.91 | 0.65 | 3.47 | 8.13 | 0.49 | 1.86
| Analytical results for V in OREAS 461 (Certified Value 385 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 382 | 382 | 390 | 392 | 398 | 411 | 375 | 370 | 400 | 362.4 | 372 | 400 | 372 |  |  | 1
| 1 | 2 | 381 | 420 | 373 | 379 | 393 | 405 | 379 | 383 | 400 | 360.2 | 374 | 400 | 368 |  |  | 6
| 1 | 3 | 384 | 394 | 379 | 392 | 413 | 397 | 379 | 355 | 400 | 363.5 | 380 | 400 | 376 |  |  | 16
| 1 | 4 | 374 | 392 | 386 | 389 | 402 | 412 | 372 | 354 | 400 | 370.3 | 370 | 400 | 372 |  |  | 384.6194
| 1 | 5 | 378 | 411 | 398 | 393 | 394 | 404 | 376 | 344 | 400 | 364.7 | 355 | 400 | 374 |  |  | 61
| 1 | 6 | 375 | 397 | 386 | 395 | 403 | 400 | 363 | 364 | 400 | 365.2 | 373 | 350 | 378
| Mean |  | 379 | 399.3333 | 385.3333 | 390 | 400.5 | 404.8333 | 374 | 361.6667 | 400 | 364.3833 | 370.6667 | 391.6667 | 373.3333
| Median |  | 379.5 | 395.5 | 386 | 392 | 400 | 404.5 | 375.5 | 359.5 | 400 | 364.1 | 372.5 | 400 | 373
| Std Dev. |  | 4 | 13.7937 | 8.6641 | 5.7271 | 7.3417 | 5.9133 | 6 | 13.7502 | 6.2269e-14 | 3.4055 | 8.3825 | 20.4124 | 3.5024
| Rel.Std.Dev. |  | 1.055% | 3.454% | 2.248% | 1.468% | 1.833% | 1.461% | 1.604% | 3.802% | 1.557e-14% | 0.9346% | 2.261% | 5.212% | 0.9381%
| PDM3 |  | -1.461% | 3.826% | 0.1856% | 1.399% | 4.129% | 5.256% | -2.761% | -5.968% | 3.999% | -5.261% | -3.628% | 1.832% | -2.934%
| Z-Score (Absolute) |  | 0.36 | 0.79 | 0 | 0.39 | 0.85 | 1.1 | 0.64 | 1.33 | 0.82 | 1.18 | 0.65 | 0.36 | 0.67
| Analytical results for W in OREAS 461 (Certified Value 3.88 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 17 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 4 | 4 | 4 | 3 | 4 | 4 | 3 | < 10 | 3 | < 5 | 4 |  |  |  |  | 1
| 1 | 2 | 4 | 4 | 4 | 3 | 4 | 4 | 3 | < 10 | 3 | < 5 | 2 |  |  |  |  | 7
| 1 | 3 | 5 | 4 | 4 | 4 | 4 | 4 | 3 | 20 | 3 | < 5 | 11 |  |  |  |  | 16
| 1 | 4 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | < 10 | 2 | < 5 | 10 |  |  |  |  | 3.8778
| 1 | 5 | 7 | 5 | 5 | 3 | 4 | 4 | 3 | < 10 | 3 | 5 | 8 |  |  |  |  | 62
| 1 | 6 | 5 | 4 | 4 | 4 | 4 | 3 | 3 | 10 | 3 | < 5 | 9
| Mean |  | 4.8333 | 4.1667 | 4.1667 | 3.5 | 4 | 3.8333 | 3 | 15 | 2.8333 | 5 | 7.3333
| Median |  | 4.5 | 4 | 4 | 3.5 | 4 | 4 | 3 | 15 | 3 | 5 | 8.5
| Std Dev. |  | 1.169 | 0.40825 | 0.40825 | 0.54772 | 9.7295e-16 | 0.40825 | 4.8648e-16 | 7.0711 | 0.40825 |  | 3.559
| Rel.Std.Dev. |  | 24.19% | 9.798% | 9.798% | 15.65% | 2.432e-14% | 10.65% | 1.622e-14% | 47.14% | 14.41% |  | 48.53%
| PDM3 |  | 24.64% | 7.45% | 7.45% | -9.742% | 3.152% | -1.146% | -22.64% | 286.8% | -26.93% | 28.94% | 89.11%
| Z-Score (Absolute) |  | 0.24 | 0 | 0 | 0.67 | 0.17 | 0.34 | 1.18 | 10.96 | 1.35 | 0.84 | 3.2
| Analytical results for Y in OREAS 461 (Certified Value 91 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 0
1
| 1 | 1 | 91.2 | 94.4 | 94 | 92.5 | 85.9 | 88.2259 | 96 | 92.3 | 73.7 | 85.5 | 79.1455 | 153 | 91 | 88 |  | 1
| 1 | 2 | 89.1 | 102 | 96 | 89.9 | 86 | 89.4633 | 98.3 | 92 | 72.3 | 86 | 82.5399 | 148 | 91 | 87.3 |  | 8
| 1 | 3 | 93 | 97.3 | 94 | 92.6 | 88.1 | 89.9939 | 94.6 | 92.3 | 70.35 | 88 | 80.3404 | 146 | 92 | 73.4 |  | 16
| 1 | 4 | 89.3 | 96.7 | 98 | 92.7 | 87.4 | 90.2965 | 94.8 | 92.9 | 71.6 | 86.5 | 82.9414 | 124 | 90 | 88.3 |  | 90.9849
| 1 | 5 | 89.3 | 101 | 96 | 92.9 | 86.9 | 90.0518 | 95.7 | 91 | 71.37 | 87 | 85.9977 | 153 | 94 | 89.5 |  | 63
| 1 | 6 | 89.5 | 95.9 | 96 | 93.3 | 87.2 | 90.5575 | 92.7 | 92.3 | 70.74 | 91 | 90.6364 | 147 | 94 | 90
| Mean |  | 90.2333 | 97.8833 | 95.6667 | 92.3167 | 86.9167 | 89.7648 | 95.35 | 92.1333 | 71.6767 | 87.3333 | 83.6002 | 145.1667 | 92 | 86.0833
| Median |  | 89.4 | 97 | 96 | 92.65 | 87.05 | 90.0228 | 95.25 | 92.3 | 71.485 | 86.75 | 82.7407 | 147.5 | 91.5 | 88.15
| Std Dev. |  | 1.5591 | 2.9822 | 1.5055 | 1.2172 | 0.84715 | 0.83727 | 1.8512 | 0.62823 | 1.2016 | 1.9916 | 4.1781 | 10.7966 | 1.6733 | 6.292
| Rel.Std.Dev. |  | 1.728% | 3.047% | 1.574% | 1.319% | 0.9747% | 0.9327% | 1.941% | 0.6819% | 1.676% | 2.281% | 4.998% | 7.437% | 1.819% | 7.309%
| PDM3 |  | -0.826% | 7.582% | 5.146% | 1.464% | -4.471% | -1.341% | 4.798% | 1.262% | -21.22% | -4.013% | -8.116% | 59.55% | 1.116% | -5.387%
| Z-Score (Absolute) |  | 0.15 | 1.14 | 0.77 | 0.2 | 0.71 | 0.23 | 0.72 | 0.17 | 3.28 | 0.64 | 1.27 | 9.85 | 0.15 | 0.42
| Analytical results for Yb in OREAS 461 (Certified Value 4.39 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.13 | 4.28 | 4.63 | 4.11 | 4.33 | 4.4039 | 4.6 | 4.4 | 3.5952 | 3.5 | 4.1 | 4.274 | 5.4 | 4 | 4.8 | 1
| 1 | 2 | 3.88 | 4.85 | 4.68 | 4.22 | 4.08 | 4.287 | 4.7 | 4.5 | 3.3622 | 3.5 | 4.2 | 4.2558 | 5.2 | 4.5 | 4.7 | 9
| 1 | 3 | 4.48 | 4.68 | 4.58 | 4.07 | 4.22 | 4.0588 | 4.5 | 4.3 | 3.5501 | 3.3 | 4.5 | 4.152 | 5 | 5 | 5.6 | 16
| 1 | 4 | 4.21 | 4.31 | 4.31 | 4.3 | 4.32 | 4.2476 | 4.5 | 4.6 | 3.6553 | 3.3 | 4.4 | 4.129 | 4.4 | 4.5 | 5 | 4.3941
| 1 | 5 | 4.14 | 4.39 | 4.8 | 4.24 | 4.42 | 4.3327 | 4.6 | 4.5 | 3.768 | 3.4 | 4.5 | 3.9089 | 5.5 | 4.5 | 4.5 | 64
| 1 | 6 | 4.13 | 4.56 | 4.16 | 4.34 | 4.18 | 4.2911 | 4.5 | 4.6 | 3.5163 | 3.3 | 4.4 | 4.05 | 5.2 | 4.5 | 4.9
| Mean |  | 4.1617 | 4.5117 | 4.5267 | 4.2133 | 4.2583 | 4.2702 | 4.5667 | 4.4833 | 3.5745 | 3.3833 | 4.35 | 4.1283 | 5.1167 | 4.5 | 4.9167
| Median |  | 4.135 | 4.475 | 4.605 | 4.23 | 4.27 | 4.2891 | 4.55 | 4.5 | 3.5727 | 3.35 | 4.4 | 4.1405 | 5.2 | 4.5 | 4.85
| Std Dev. |  | 0.1926 | 0.22551 | 0.24213 | 0.10539 | 0.12205 | 0.11638 | 0.08165 | 0.1169 | 0.13671 | 0.098319 | 0.16432 | 0.13587 | 0.392 | 0.31623 | 0.37639
| Rel.Std.Dev. |  | 4.628% | 4.998% | 5.349% | 2.501% | 2.866% | 2.725% | 1.788% | 2.608% | 3.825% | 2.906% | 3.777% | 3.291% | 7.661% | 7.027% | 7.655%
| PDM3 |  | -5.289% | 2.676% | 3.018% | -4.113% | -3.089% | -2.819% | 3.928% | 2.032% | -18.65% | -23% | -1.003% | -6.049% | 16.45% | 2.411% | 11.89%
| Z-Score (Absolute) |  | 0.96 | 0.62 | 0.67 | 0.52 | 0.35 | 0.14 | 0.83 | 0.51 | 2.96 | 3.69 | 0 | 0.85 | 3.47 | 0.57 | 1.64
| Analytical results for Zn in OREAS 461 (Certified Value 162 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 14 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 143 | 200 | 242.3984 | 190 | 150 | 138 | [9 empty] | 1
| 1 | 2 | 152 | 200 | 251.6671 | 170 | 150 | 132 | [9 empty] | #N/A
| 1 | 3 | 146 | 200 | 256.3602 | 170 | 150 | 138 | [9 empty] | 16
| 1 | 4 | 153 | 200 | 247.1013 | 170 | 150 | 139 | [9 empty] | 161.7333
| 1 | 5 | 156 | 200 | 234.7434 | 180 | 150 | 132 | [9 empty] | 65
| 1 | 6 | 158 | 200 | 259.2983 | 170 | 150 | 133
| Mean |  | 151.3333 | 200 | 248.5948 | 175 | 150 | 135.3333
| Median |  | 152.5 | 200 | 249.3842 | 170 | 150 | 135.5
| Std Dev. |  | 5.785 | 3.1134e-14 | 9.1301 | 8.3666 | 3.1134e-14 | 3.3267
| Rel.Std.Dev. |  | 3.823% | 1.557e-14% | 3.673% | 4.781% | 2.076e-14% | 2.458%
| PDM3 |  | -6.43% | 23.66% | 53.71% | 8.203% | -7.255% | -16.32%
| Z-Score (Absolute) |  | 0.37 | 1.36 | 3.09 | 0.37 | 0.41 | 0.93
| Analytical results for Zr in OREAS 461 (Certified Value 603 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 596 | 631 | 614 | 615 | 585 | 619 | 600 | 587 | 561 | 580.4 | 658 |  |  |  |  | 1
| 1 | 2 | 604 | 683 | 625 | 608 | 599 | 621 | 579 | 625 | 576 | 586.3 | 657 |  |  |  |  | 10
| 1 | 3 | 607 | 652 | 627 | 610 | 605 | 613 | 582 | 584 | 569 | 564.1 | 637 |  |  |  |  | 16
| 1 | 4 | 587 | 637 | 635 | 627 | 624 | 609 | 588 | 620 | 565 | 581.1 | 537 |  |  |  |  | 603.0104
| 1 | 5 | 573 | 658 | 651 | 608 | 617 | 609 | 588 | 585 | 552 | 567.8 | 680 |  |  |  |  | 66
| 1 | 6 | 581 | 623 | 634 | 629 | 594 | 609 | 602 | 589 | 565 | 564.9 | 651
| Mean |  | 591.3333 | 647.3333 | 631 | 616.1667 | 604 | 613.3333 | 589.8333 | 598.3333 | 564.6667 | 574.1 | 636.6667
| Median |  | 591.5 | 644.5 | 630.5 | 612.5 | 602 | 611 | 588 | 588 | 565 | 574.1 | 654
| Std Dev. |  | 13.3367 | 21.7868 | 12.3774 | 9.5376 | 14.5327 | 5.4283 | 9.347 | 18.8644 | 8.0166 | 9.611 | 50.7648
| Rel.Std.Dev. |  | 2.255% | 3.366% | 1.962% | 1.548% | 2.406% | 0.8851% | 1.585% | 3.153% | 1.42% | 1.674% | 7.974%
| PDM3 |  | -1.936% | 7.35% | 4.642% | 2.182% | 0.1641% | 1.712% | -2.185% | -0.7756% | -6.359% | -4.794% | 5.581%
| Z-Score (Absolute) |  | 0.6 | 2.06 | 1.29 | 0.58 | 0 | 0.44 | 0.67 | 0.27 | 1.87 | 1.42 | 2.5
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 461 (Certified Value < 2 ppm) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 0.59 | 1.01 | < 0.01 | 0.18 | 0.26 | 0.47462 | 0.6 | 0.9 | < 0.01 | 1.46 | < 1 | 1 |  |  | 1
| 1 | 2 | 0.58 | 1.14 | 0.86 | 0.22 | 0.21 | 0.44091 | 0.8 | 0.9 | < 0.01 | 1.02 | < 1 | 0.5 |  |  | 15
| 1 | 3 | 0.33 | 1.07 | < 0.01 | 0.13 | 0.33 | 0.41447 | 0.7 | 1 | < 0.01 | 1.56 | < 1 | 0.5 |  |  | 16
| 1 | 4 | 0.71 | 1.2 | 1.86 | < 0.01 | 0.33 | 0.4622 | 0.6 | 0.7 | < 0.01 | 1.43 | < 1 | 1 |  |  | < 2
| 1 | 5 | 0.09 | 0.91 | < 0.01 | 0.48 | 0.38 | 0.50033 | 0.9 | 1.5 | < 0.01 | 1.25 | < 1 | 1 |  |  | 68
| 1 | 6 | 0.96 | 1 | 1.4 | 0.09 | 0.24 | 0.50642 | 0.7 | 1.3 | < 0.01 | 1.49 | < 1 | 1
| Mean |  | 0.54333 | 1.055 | 1.3733 | 0.22 | 0.29167 | 0.46649 | 0.71667 | 1.05 |  | 1.3683 |  | 0.83333
| Median |  | 0.585 | 1.04 | 1.4 | 0.18 | 0.295 | 0.46841 | 0.7 | 0.95 |  | 1.445 |  | 1
| Std Dev. |  | 0.30197 | 0.10445 | 0.50053 | 0.15346 | 0.064936 | 0.035169 | 0.1169 | 0.29496 |  | 0.19954 |  | 0.2582
| Rel.Std.Dev. |  | 55.58% | 9.901% | 36.45% | 69.75% | 22.26% | 7.539% | 16.31% | 28.09% |  | 14.58% |  | 30.98%
PDM3
| Z-Score (Absolute) |  | 0.53 | 0.64 | 1.37 | 1.42 | 1.11 | 0.71 | 0.13 | 0.63 | NA | 1.52 | NA | 0.13
| Analytical results for Al in OREAS 461 (Certified Value 5.96 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 5.11 | 5.98 | 5.73 | 6.22 | 6.15 | 5.7075 | 5.95 | 5.86 | 6.02 | 5.8334 | 5.9812 | 5.94 | 6.13 | 6.19 | 1
| 1 | 2 | 5.69 | 6.05 | 5.73 | 6.15 | 6.09 | 5.5674 | 5.87 | 5.84 | 6.01 | 5.9219 | 6.0334 | 5.72 | 6.142 | 6.17 | 1
| 1 | 3 | 5.36 | 6.1 | 5.71 | 6.23 | 5.79 | 5.6788 | 5.85 | 5.94 | 6.01 | 5.8959 | 5.9544 | 5.85 | 6.119 | 6.27 | 16
| 1 | 4 | 5.52 | 6.07 | 5.76 | 6.18 | 5.64 | 5.937 | 5.93 | 5.88 | 6.04 | 6.0783 | 5.9523 | 5.99 | 6.124 | 6.09 | 5.9585
| 1 | 5 | 5.03 | 6.1 | 5.72 | 6.33 | 5.97 | 5.7769 | 5.91 | 5.91 | 5.97 | 5.8137 | 5.8569 | 5.83 | 6.177 | 6.16 | 69
| 1 | 6 | 4.94 | 6.13 | 5.72 | 6.29 | 5.64 | 5.6531 | 5.91 | 6.01 | 5.99 | 5.8489 | 5.854 | 5.86 | 6.115 | 6.38
| Mean |  | 5.275 | 6.0717 | 5.7283 | 6.2333 | 5.88 | 5.7201 | 5.9033 | 5.9067 | 6.0067 | 5.8987 | 5.9387 | 5.865 | 6.1345 | 6.21
| Median |  | 5.235 | 6.085 | 5.725 | 6.225 | 5.88 | 5.6932 | 5.91 | 5.895 | 6.01 | 5.8724 | 5.9533 | 5.855 | 6.127 | 6.18
| Std Dev. |  | 0.29629 | 0.052694 | 0.017224 | 0.067132 | 0.22289 | 0.12643 | 0.037238 | 0.06186 | 0.024221 | 0.096717 | 0.070804 | 0.093541 | 0.022863 | 0.10139
| Rel.Std.Dev. |  | 5.617% | 0.8679% | 0.3007% | 1.077% | 3.791% | 2.21% | 0.6308% | 1.047% | 0.4032% | 1.64% | 1.192% | 1.595% | 0.3727% | 1.633%
| PDM3 |  | -11.47% | 1.898% | -3.864% | 4.612% | -1.318% | -4.001% | -0.9266% | -0.8707% | 0.8076% | -1.005% | -0.3329% | -1.57% | 2.953% | 4.22%
| Z-Score (Absolute) |  | 3.17 | 0.84 | 0.89 | 1.65 | 0.13 | 0.93 | 0.01 | 0.01 | 0.51 | 0.21 | 0.17 | 0.2 | 1.15 | 1.53
| Analytical results for As in OREAS 461 (Certified Value 33.6 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 34 | 34.8 | 34.2 | 33 | 33.2 | 30.5833 | 34 | 34 | 33 | 48.2 | < 50 | 35 |  |  | 1
| 1 | 2 | 33.5 | 32.5 | 34.6 | 32.7 | 31.9 | 27.5692 | 36 | 32 | 33.2 | 44.1 | < 50 | 34 |  |  | 16
| 1 | 3 | 33.8 | 33.4 | 35.4 | 32.5 | 33 | 28.2683 | 35 | 35 | 32.5 | 49.7 | < 50 | 37 |  |  | 16
| 1 | 4 | 34.8 | 33.3 | 32.3 | 32.5 | 32 | 30.45 | 35 | 34 | 32.4 | 51.1 | < 50 | 34 |  |  | 33.5537
| 1 | 5 | 34 | 32.8 | 34.2 | 32.8 | 33.1 | 29.5769 | 33 | 33 | 32 | 50.4 | < 50 | 35 |  |  | 70
| 1 | 6 | 33.9 | 32.3 | 32.6 | 32.2 | 33.2 | 29.8258 | 33 | 36 | 32.7 | 50.5 | < 50 | 35
| Mean |  | 34 | 33.1833 | 33.8833 | 32.6167 | 32.7333 | 29.3789 | 34.3333 | 34 | 32.6333 | 49 |  | 35
| Median |  | 33.95 | 33.05 | 34.2 | 32.6 | 33.05 | 29.7014 | 34.5 | 34 | 32.6 | 50.05 |  | 35
| Std Dev. |  | 0.43359 | 0.90203 | 1.1974 | 0.27869 | 0.6121 | 1.212 | 1.2111 | 1.4142 | 0.43205 | 2.5985 |  | 1.0954
| Rel.Std.Dev. |  | 1.275% | 2.718% | 3.534% | 0.8544% | 1.87% | 4.126% | 3.527% | 4.159% | 1.324% | 5.303% |  | 3.13%
| PDM3 |  | 1.33% | -1.104% | 0.9824% | -2.793% | -2.445% | -12.44% | 2.324% | 1.33% | -2.743% | 46.03% |  | 4.31%
| Z-Score (Absolute) |  | 0.11 | 0.66 | 0 | 1.19 | 1.08 | 4.24 | 0.42 | 0.11 | 1.18 | 15.15 | NA | 0.67
| Analytical results for Ba in OREAS 461 (Certified Value 957 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 930 | 950 | 980 | 1020 | 1000 | 930.1279 | 937 | 913 | 1020 | 945.2 | 1030.99 | 330 | 945 | 918 | 1
| 1 | 2 | 950 | 950 | 970 | 1000 | 1000 | 926.4228 | 945 | 928 | 1040 | 952.3 | 1026.01 | 334 | 955 | 890 | 2
| 1 | 3 | 880 | 950 | 970 | 1030 | 950 | 940.008 | 970 | 916 | 1050 | 952.1 | 1005.06 | 348 | 950 | 896 | 16
| 1 | 4 | 890 | 940 | 960 | 1010 | 940 | 926.149 | 945 | 911 | 1030 | 968.7 | 1018.382 | 335 | 955 | 898 | 957.0984
| 1 | 5 | 870 | 940 | 960 | 1030 | 1000 | 949.4869 | 931 | 921 | 1020 | 922.5 | 1009.25 | 350 | 945 | 899 | 71
| 1 | 6 | 880 | 950 | 970 | 1020 | 920 | 932.7722 | 932 | 932 | 1040 | 931.1 | 1004.92 | 352 | 940 | 950
| Mean |  | 900 | 946.6667 | 968.3333 | 1018.3333 | 968.3333 | 934.1611 | 943.3333 | 920.1667 | 1033.3333 | 945.3167 | 1015.7687 | 341.5 | 948.3333 | 908.5
| Median |  | 885 | 950 | 970 | 1020 | 975 | 931.45 | 941 | 918.5 | 1035 | 948.65 | 1013.816 | 341.5 | 947.5 | 898.5
| Std Dev. |  | 32.249 | 5.164 | 7.5277 | 11.6905 | 36.0093 | 9.0701 | 14.4037 | 8.4242 | 12.1106 | 16.5184 | 11.1199 | 9.5446 | 6.0553 | 22.4121
| Rel.Std.Dev. |  | 3.583% | 0.5455% | 0.7774% | 1.148% | 3.719% | 0.9709% | 1.527% | 0.9155% | 1.172% | 1.747% | 1.095% | 2.795% | 0.6385% | 2.467%
| PDM3 |  | -5.966% | -1.09% | 1.174% | 6.398% | 1.174% | -2.397% | -1.438% | -3.859% | 7.965% | -1.231% | 6.13% | -64.32% | -0.9158% | -5.078%
| Z-Score (Absolute) |  | 1.29 | 0.02 | 0.63 | 2.03 | 0.63 | 0.33 | 0.07 | 0.72 | 2.45 | 0.02 | 1.95 | 16.93 | 0.07 | 1.28
| Analytical results for Be in OREAS 461 (Certified Value 2.09 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 1.6 | 2.36 | 1.73 | 2.4 | 1.99 | 1.8332 | 2.3 | 2.4 | 2.04 | 2.16 | 1.9561 | 3 | 2.3 |  | 1
| 1 | 2 | 1.65 | 2.36 | 1.9 | 2.24 | 1.92 | 1.9227 | 2.3 | 1.5 | 2.01 | 2.26 | 2.2213 | 3 | 2.2 |  | #N/A
| 1 | 3 | 1.66 | 2.32 | 1.79 | 2.32 | 2.03 | 2.0916 | 2.4 | 1.9 | 2.04 | 2.33 | 2.0178 | 2.5 | 2.3 |  | 16
| 1 | 4 | 1.66 | 2.31 | 2.01 | 2.44 | 1.87 | 2.1006 | 2.3 | 1.9 | 2.02 | 2.32 | 2.0433 | 2.5 | 2.1 |  | 2.0913
| 1 | 5 | 1.77 | 2.14 | 1.83 | 2.36 | 1.9 | 2.1986 | 2.3 | 2.7 | 2 | 2.24 | 1.9839 | 3 | 2.1 |  | 72
| 1 | 6 | 1.72 | 2.14 | 1.83 | 2.5 | 1.87 | 2.2449 | 2.2 | 2.2 | 2.02 | 2.34 | 2.0687 | 2.5 | 2.3
| Mean |  | 1.6767 | 2.2717 | 1.8483 | 2.3767 | 1.93 | 2.0652 | 2.3 | 2.1 | 2.0217 | 2.275 | 2.0485 | 2.75 | 2.2167
| Median |  | 1.66 | 2.315 | 1.83 | 2.38 | 1.91 | 2.0961 | 2.3 | 2.05 | 2.02 | 2.29 | 2.0305 | 2.75 | 2.25
| Std Dev. |  | 0.059554 | 0.104 | 0.096833 | 0.091579 | 0.06603 | 0.15885 | 0.063246 | 0.42426 | 0.016021 | 0.06921 | 0.093758 | 0.27386 | 0.098319
| Rel.Std.Dev. |  | 3.552% | 4.578% | 5.239% | 3.853% | 3.421% | 7.692% | 2.75% | 20.2% | 0.7925% | 3.042% | 4.577% | 9.959% | 4.435%
| PDM3 |  | -19.83% | 8.623% | -11.62% | 13.64% | -7.714% | -1.247% | 9.978% | 0.4149% | -3.331% | 8.783% | -2.047% | 31.5% | 5.994%
| Z-Score (Absolute) |  | 1.66 | 0.67 | 0.99 | 1.09 | 0.67 | 0.14 | 0.79 | 0 | 0.31 | 0.69 | 0.34 | 2.55 | 0.46
| Analytical results for Bi in OREAS 461 (Certified Value 2.43 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.55 | 2.48 | 2.54 | 2.4 | 2.18 | 2.2786 | 2.5 | 2.6 | < 0.01 | 2.48 | 2.3108 | 2.4 | 2.3 |  | 1
| 1 | 2 | 2.56 | 2.48 | 2.58 | 2.39 | 2.14 | 2.1842 | 2.5 | 2.7 | < 0.01 | 2.49 | 2.0975 | 3.8 | 2.3 |  | #N/A
| 1 | 3 | 2.57 | 2.48 | 2.63 | 2.47 | 2.22 | 2.3718 | 2.6 | 2.6 | < 0.01 | 2.55 | 2.2453 | 2.2 | 2.4 |  | 16
| 1 | 4 | 2.58 | 2.55 | 2.57 | 2.52 | 2.12 | 2.4188 | 2.6 | 2.5 | < 0.01 | 2.58 | 2.0735 | 2.4 | 2.3 |  | 2.4269
| 1 | 5 | 2.51 | 2.53 | 2.62 | 2.48 | 2.17 | 2.4444 | 2.6 | 2.6 | < 0.01 | 2.46 | 2.1496 | 2.4 | 2.4 |  | 73
| 1 | 6 | 2.57 | 2.45 | 2.49 | 2.51 | 2.19 | 2.2143 | 2.4 | 2.7 | < 0.01 | 2.42 | 2.1767 | 2.4 | 2.5
| Mean |  | 2.5567 | 2.495 | 2.5717 | 2.4617 | 2.17 | 2.3187 | 2.5333 | 2.6167 |  | 2.4967 | 2.1756 | 2.6 | 2.3667
| Median |  | 2.565 | 2.48 | 2.575 | 2.475 | 2.175 | 2.3252 | 2.55 | 2.6 |  | 2.485 | 2.1632 | 2.4 | 2.35
| Std Dev. |  | 0.025033 | 0.037283 | 0.051929 | 0.054924 | 0.035777 | 0.10886 | 0.08165 | 0.075277 |  | 0.058878 | 0.089841 | 0.5933 | 0.08165
| Rel.Std.Dev. |  | 0.9791% | 1.494% | 2.019% | 2.231% | 1.649% | 4.695% | 3.223% | 2.877% |  | 2.358% | 4.129% | 22.82% | 3.45%
| PDM3 |  | 5.348% | 2.807% | 5.966% | 1.433% | -10.58% | -4.458% | 4.386% | 7.82% |  | 2.875% | -10.35% | 7.133% | -2.481%
| Z-Score (Absolute) |  | 0.52 | 0.11 | 0.61 | 0.11 | 2.03 | 1.05 | 0.36 | 0.91 | NA | 0.12 | 1.99 | 0.78 | 0.73
| Analytical results for Ca in OREAS 461 (Certified Value 1.21 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.23 | 1.25 | 1.06 | 1.14 | 1.07 | 1.155 | 1.22 | 1.29 | 1.16 | 1.3214 | 1.2546 | 1.18 | 1.2735 | 1.37 | 1
| 1 | 2 | 1.25 | 1.24 | 1.06 | 1.13 | 1.11 | 1.158 | 1.21 | 1.29 | 1.17 | 1.3087 | 1.2626 | 1.17 | 1.2725 | 1.37 | 3
| 1 | 3 | 1.21 | 1.26 | 1.04 | 1.15 | 1.07 | 1.1894 | 1.2 | 1.31 | 1.17 | 1.3334 | 1.254 | 1.2 | 1.272 | 1.4 | 16
| 1 | 4 | 1.23 | 1.29 | 1.06 | 1.14 | 1.02 | 1.2153 | 1.22 | 1.28 | 1.17 | 1.3683 | 1.2406 | 1.19 | 1.2715 | 1.37 | 1.2147
| 1 | 5 | 1.2 | 1.26 | 1.06 | 1.18 | 1.1 | 1.1753 | 1.22 | 1.29 | 1.16 | 1.268 | 1.2505 | 1.15 | 1.2675 | 1.38 | 74
| 1 | 6 | 1.21 | 1.26 | 1.06 | 1.16 | 1.02 | 1.1478 | 1.22 | 1.32 | 1.16 | 1.3018 | 1.2446 | 1.2 | 1.2685 | 1.4
| Mean |  | 1.2217 | 1.26 | 1.0567 | 1.15 | 1.065 | 1.1735 | 1.215 | 1.2967 | 1.165 | 1.3169 | 1.2511 | 1.1817 | 1.2709 | 1.3817
| Median |  | 1.22 | 1.26 | 1.06 | 1.145 | 1.07 | 1.1667 | 1.22 | 1.29 | 1.165 | 1.3151 | 1.2522 | 1.185 | 1.2717 | 1.375
| Std Dev. |  | 0.018348 | 0.016733 | 0.008165 | 0.017889 | 0.038341 | 0.025496 | 0.0083666 | 0.015055 | 0.0054772 | 0.033529 | 0.0078239 | 0.019408 | 0.0023752 | 0.01472
| Rel.Std.Dev. |  | 1.502% | 1.328% | 0.7727% | 1.556% | 3.6% | 2.173% | 0.6886% | 1.161% | 0.4701% | 2.546% | 0.6253% | 1.642% | 0.1869% | 1.065%
| PDM3 |  | 0.5737% | 3.729% | -13.01% | -5.326% | -12.32% | -3.395% | 0.02482% | 6.748% | -4.091% | 8.416% | 3% | -2.719% | 4.628% | 13.75%
| Z-Score (Absolute) |  | 0.04 | 0.53 | 2.06 | 0.87 | 1.95 | 0.57 | 0.04 | 1 | 0.68 | 1.26 | 0.42 | 0.47 | 0.67 | 2.08
| Analytical results for Cd in OREAS 461 (Certified Value < 0.1 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.04 | 0.05 | 0.07 | 0.04 | 0.07 | 0.07 | < 0.2 | 0.44 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 1
| 1 | 2 | 0.06 | < 0.02 | 0.08 | 0.07 | 0.05 | 0.076464 | < 0.2 | 0.41 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 18
| 1 | 3 | 0.08 | 0.04 | 0.06 | < 0.02 | 0.05 | 0.07 | < 0.2 | 0.4 | 0.03 | < 0.5 | < 0.5 |  |  |  | 16
| 1 | 4 | 0.04 | 0.03 | 0.09 | 0.05 | 0.05 | 0.07 | < 0.2 | 0.41 | < 0.02 | < 0.5 | < 0.5 |  |  |  | < 0.1
| 1 | 5 | 0.03 | 0.04 | 0.04 | 0.03 | 0.06 | 0.07 | 0.2 | 0.4 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 75
| 1 | 6 | 0.05 | 0.06 | 0.03 | 0.04 | 0.06 | 0.071069 | < 0.2 | 0.39 | < 0.02 | < 0.5 | < 0.5
| Mean |  | 0.05 | 0.044 | 0.061667 | 0.046 | 0.056667 | 0.071256 | 0.2 | 0.40833 | 0.03
| Median |  | 0.045 | 0.04 | 0.065 | 0.04 | 0.055 | 0.07 | 0.2 | 0.405 | 0.03
| Std Dev. |  | 01: 0.017889; 02: 0.011402; 03: 0.023166; 05: 0.015166; 06: 0.008165; 07: 0.0025873; 09: 0.017224
| Rel.Std.Dev. |  | 01: 35.78%; 02: 25.91%; 03: 37.57%; 05: 32.97%; 06: 14.41%; 07: 3.631%; 09: 4.218%
PDM3
| Z-Score (Absolute) |  | 0.27 | 0.74 | 0.2 | 0.67 | 0 | 0.59 | 2.43 | 14.23 | 1.75 | NA | NA
| Analytical results for Ce in OREAS 461 (Certified Value 3508 ppm) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 3526.9517 | 3428.7 | 3266 | 3480 | 3289.37 | 3580.4551 | 3696 | 3440 |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 3587.9826 | 3441 | 3457.7 | 3520 | 3396.74 | 3625.3739 | 3628 | 3470 |  | #N/A
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 3548.9624 | 3511.6 | 3372 | 3500 | 3440.42 | 3623.758 | 3682 | 3520 |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 3572.1857 | 3527.9 | 3364.9 | 3590 | 3406.87 | 3557.8769 | 3678.5 | 3450 |  | 3508.4834
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 3479.4713 | 3440.2 | 3435.9 | 3560 | 3372.8 | 3586.7322 | 3679.5 | 3520 |  | 76
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 3535.18 | 3436.4 | 3448.6 | 3460 | 3268.71 | 3586.461 | 3766 | 3650
| Mean | [6 empty] | 3541.789 | 3464.3 | 3390.85 | 3518.3333 | 3362.485 | 3593.4429 | 3688.3333 | 3508.3333
| Median | [6 empty] | 3542.0712 | 3440.6 | 3403.95 | 3510 | 3384.77 | 3586.5966 | 3680.75 | 3495
| Std Dev. | [6 empty] | 38.0962 | 43.4781 | 72.732 | 49.1596 | 68.5029 | 26.3287 | 44.6001 | 77.3089
| Rel.Std.Dev. | [6 empty] | 1.076% | 1.255% | 2.145% | 1.397% | 2.037% | 0.7327% | 1.209% | 2.204%
| PDM3 | [6 empty] | 0.9493% | -1.259% | -3.353% | 0.2807% | -4.161% | 2.422% | 5.126% | -0.004276%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.3 | 0.51 | 1.28 | 0.05 | 1.58 | 0.84 | 1.83 | 0.05
| Analytical results for Co in OREAS 461 (Certified Value 12.1 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 11.6 | 12 | 11.8 | 12.4 | 12 | 10.5852 | 13 | 13 | 11.7 | 13.1 | 12.0447 | 25 | 15 |  | 1
| 1 | 2 | 11.9 | 11.7 | 11.9 | 11.9 | 11.5 | 10.638 | 13 | 13 | 11.7 | 12.8 | 11.9276 | 25 | 10 |  | #N/A
| 1 | 3 | 11.9 | 12 | 12.2 | 12.3 | 12 | 10.4673 | 13.3 | 13 | 12 | 12.7 | 11.8657 | 25 | 15 |  | 16
| 1 | 4 | 11.9 | 12.2 | 11.5 | 12.1 | 11.4 | 11.3589 | 13 | 13 | 11.9 | 12.9 | 11.9513 | 25 | 15 |  | 12.1248
| 1 | 5 | 12 | 11.7 | 12.1 | 12.3 | 11.6 | 11.2319 | 12.9 | 13 | 11.8 | 12.7 | 12.126 | 25 | 10 |  | 77
| 1 | 6 | 11.8 | 12 | 12 | 12.5 | 11.8 | 11.4416 | 12.9 | 12 | 12.1 | 12.4 | 12.0241 | 25 | 15
| Mean |  | 11.85 | 11.9333 | 11.9167 | 12.25 | 11.7167 | 10.9538 | 13.0167 | 12.8333 | 11.8667 | 12.7667 | 11.9899 | 25 | 13.3333
| Median |  | 11.9 | 12 | 11.95 | 12.3 | 11.7 | 10.935 | 13 | 13 | 11.85 | 12.75 | 11.9877 | 25 | 15
| Std Dev. |  | 0.13784 | 0.19664 | 0.24833 | 0.21679 | 0.25626 | 0.43627 | 0.1472 | 0.40825 | 0.1633 | 0.23381 | 0.093297 | 0 | 2.582
| Rel.Std.Dev. |  | 1.163% | 1.648% | 2.084% | 1.77% | 2.187% | 3.983% | 1.131% | 3.181% | 1.376% | 1.831% | 0.7781% | 0% | 19.36%
| PDM3 |  | -2.266% | -1.579% | -1.717% | 1.033% | -3.366% | -9.658% | 7.356% | 5.844% | -2.129% | 5.294% | -1.113% | 106.2% | 9.967%
| Z-Score (Absolute) |  | 0.35 | 0.14 | 0.18 | 0.64 | 0.67 | 2.56 | 2.53 | 2.08 | 0.3 | 1.92 | 0 | 32.11 | 3.32
| Analytical results for Cr in OREAS 461 (Certified Value 524 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 511 | 552 | 438 | 508 | 474 | 539.6079 | 529 | 521 | 641 | 391 | 600 |  |  |  | 1
| 1 | 2 | 510 | 544 | 440 | 501 | 496 | 542.0672 | 526 | 524 | 648 | 387 | 600 |  |  |  | 4
| 1 | 3 | 492 | 555 | 433 | 516 | 472 | 558.2945 | 571 | 514 | 637 | 421 | 600 |  |  |  | 16
| 1 | 4 | 499 | 572 | 441 | 504 | 454 | 568.1738 | 547 | 533 | 686 | 397 | 650 |  |  |  | 523.6841
| 1 | 5 | 494 | 553 | 443 | 523 | 486 | 562.6838 | 550 | 506 | 652 | 382 | 600 |  |  |  | 78
| 1 | 6 | 504 | 552 | 441 | 514 | 456 | 568.5218 | 566 | 505 | 679 | 386 | 650
| Mean |  | 501.6667 | 554.6667 | 439.3333 | 511 | 473 | 556.5582 | 548.1667 | 517.1667 | 657.1667 | 394 | 616.6667
| Median |  | 501.5 | 552.5 | 440.5 | 511 | 473 | 560.4892 | 548.5 | 517.5 | 650 | 389 | 600
| Std Dev. |  | 8.0166 | 9.288 | 3.5024 | 8.1976 | 16.4317 | 12.7741 | 18.4544 | 10.9072 | 20.4296 | 14.1704 | 25.8199
| Rel.Std.Dev. |  | 1.598% | 1.675% | 0.7972% | 1.604% | 3.474% | 2.295% | 3.367% | 2.109% | 3.109% | 3.597% | 4.187%
| PDM3 |  | -4.204% | 5.916% | -16.11% | -2.422% | -9.678% | 6.277% | 4.675% | -1.245% | 25.49% | -24.76% | 17.76%
| Z-Score (Absolute) |  | 0.27 | 0.58 | 1.33 | 0.11 | 0.76 | 0.67 | 0.53 | 0 | 2.4 | 2.2 | 1.7
| Analytical results for Cs in OREAS 461 (Certified Value 0.77 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.51 | 0.79 | 0.81 | 0.76 | 0.73 | 0.7631 | 0.8 | 0.8 | 0.84 | 0.7 | 1.0552 | 1 | 0.7 |  | 1
| 1 | 2 | 0.72 | 0.88 | 0.85 | 0.78 | 0.71 | 0.7181 | 0.8 | 0.9 | 0.85 | 0.7 | 1.043 | 1 | 0.7 |  | #N/A
| 1 | 3 | 0.74 | 0.78 | 0.86 | 0.77 | 0.73 | 0.74091 | 0.8 | 0.9 | 0.85 | 0.74 | 0.98198 | 1 | 0.7 |  | 16
| 1 | 4 | 0.8 | 0.79 | 0.79 | 0.81 | 0.71 | 0.72047 | 0.8 | 0.9 | 0.86 | 0.72 | 0.98982 | 1 | 0.7 |  | 0.76887
| 1 | 5 | 0.64 | 0.78 | 0.85 | 0.84 | 0.72 | 0.78307 | 0.8 | 0.8 | 0.84 | 0.69 | 0.99982 | 1 | 0.7 |  | 79
| 1 | 6 | 0.55 | 0.79 | 0.82 | 0.81 | 0.71 | 0.71156 | 0.8 | 0.8 | 0.84 | 0.7 | 1.0944 | 1 | 0.7
| Mean |  | 0.66 | 0.80167 | 0.83 | 0.795 | 0.71833 | 0.73954 | 0.8 | 0.85 | 0.84667 | 0.70833 | 1.0274 | 1 | 0.7
| Median |  | 0.68 | 0.79 | 0.835 | 0.795 | 0.715 | 0.73069 | 0.8 | 0.85 | 0.845 | 0.7 | 1.0214 | 1 | 0.7
| Std Dev. |  | 0.11367 | 0.038687 | 0.027568 | 0.030166 | 0.0098319 | 0.028489 | 1.2162e-16 | 0.054772 | 0.008165 | 0.018348 | 0.044132 | 0 | 1.2162e-16
| Rel.Std.Dev. |  | 17.22% | 4.826% | 3.321% | 3.794% | 1.369% | 3.852% | 1.52e-14% | 6.444% | 0.9644% | 2.59% | 4.296% | 0% | 1.737e-14%
| PDM3 |  | -14.16% | 4.266% | 7.951% | 3.399% | -6.572% | -3.815% | 4.049% | 10.55% | 10.12% | -7.873% | 33.62% | 30.06% | -8.957%
| Z-Score (Absolute) |  | 1.28 | 0.11 | 0.43 | 0 | 0.93 | 0.67 | 0.06 | 0.67 | 0.63 | 1.13 | 2.83 | 2.49 | 1.15
| Analytical results for Cu in OREAS 461 (Certified Value 56 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 56.1 | 56.6 | 56.6 | 55 | 58 | 51.9975 | 58.3 | 69 | 57.4 | 98 | 55.2444 | 26.7 | 50 | 74 | 1
| 1 | 2 | 56.4 | 55.5 | 56.2 | 53.5 | 54.5 | 51.9646 | 58.1 | 70 | 57.2 | 96 | 56.1858 | 24.2 | 60 | 72 | #N/A
| 1 | 3 | 56 | 56.7 | 57.7 | 55.9 | 56 | 53.2252 | 58.6 | 79 | 58 | 96 | 55.8862 | 28.3 | 50 | 74 | 16
| 1 | 4 | 57.4 | 56.3 | 52.1 | 54.9 | 54.1 | 54.4716 | 58.4 | 67 | 57.3 | 100 | 55.2083 | 27.4 | 50 | 74 | 55.8138
| 1 | 5 | 56.3 | 56.1 | 56.9 | 55.4 | 56 | 53.8908 | 57.6 | 64 | 57.3 | 100 | 55.4457 | 27.8 | 60 | 74 | 80
| 1 | 6 | 56.9 | 55.6 | 56.8 | 55.6 | 55.3 | 53.5101 | 58.4 | 61 | 58.4 | 85 | 55.6601 | 27.1 | 50 | 74
| Mean |  | 56.5167 | 56.1333 | 56.05 | 55.05 | 55.65 | 53.1766 | 58.2333 | 68.3333 | 57.6 | 95.8333 | 55.6051 | 26.9167 | 53.3333 | 73.6667
| Median |  | 56.35 | 56.2 | 56.7 | 55.2 | 55.65 | 53.3677 | 58.35 | 68 | 57.35 | 97 | 55.5529 | 27.25 | 50 | 74
| Std Dev. |  | 0.53448 | 0.50067 | 1.9967 | 0.84558 | 1.3867 | 1.0159 | 0.35024 | 6.186 | 0.4858 | 5.6006 | 0.38268 | 1.4414 | 5.164 | 0.8165
| Rel.Std.Dev. |  | 0.9457% | 0.8919% | 3.562% | 1.536% | 2.492% | 1.91% | 0.6014% | 9.053% | 0.8434% | 5.844% | 0.6882% | 5.355% | 9.682% | 1.108%
| PDM3 |  | 1.259% | 0.5724% | 0.4231% | -1.369% | -0.2935% | -4.725% | 4.335% | 22.43% | 3.2% | 71.7% | -0.374% | -51.77% | -4.444% | 31.99%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 0.22 | 0.54 | 0.29 | 1.33 | 0.81 | 5.09 | 0.54 | 17.66 | 0.3 | 12.23 | 1.27 | 7.35
| Analytical results for Dy in OREAS 461 (Certified Value 33.5 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 32 | 31.7 | 34.9 | 32.4 | 34.1 | 33.6238 | 33 | 32.2 | 35.1 | 33.8 | 35.9756 | 32.9 | 34.9 |  | 1
| 1 | 2 | 34.1 | 33.1 | 34.2 | 32.7 | 32.4 | 34.4217 | 32.7 | 33.3 | 35 | 33.64 | 35.6305 | 32.9 | 31.8 |  | 5
| 1 | 3 | 34 | 32.6 | 34.3 | 33.5 | 33.8 | 34.1324 | 33.4 | 33.7 | 34.6 | 33.46 | 35.9598 | 33.3 | 32.8 |  | 16
| 1 | 4 | 35.4 | 32.9 | 31.1 | 33.7 | 32.4 | 34.1261 | 33.4 | 33 | 33.8 | 34.07 | 35.5468 | 33 | 32.7 |  | 33.5055
| 1 | 5 | 33.5 | 32.4 | 34.7 | 33.6 | 33.5 | 34.0991 | 32.6 | 33.9 | 34.6 | 33.34 | 35.7708 | 33.1 | 31.7 |  | 81
| 1 | 6 | 32.9 | 32.3 | 33.5 | 34.2 | 34.1 | 34.5262 | 33 | 34.1 | 34.7 | 33.24 | 36.4661 | 33.4 | 34.1
| Mean |  | 33.65 | 32.5 | 33.7833 | 33.35 | 33.3833 | 34.1549 | 33.0167 | 33.3667 | 34.6333 | 33.5917 | 35.8916 | 33.1 | 33
| Median |  | 33.75 | 32.5 | 34.25 | 33.55 | 33.65 | 34.1292 | 33 | 33.5 | 34.65 | 33.55 | 35.8653 | 33.05 | 32.75
| Std Dev. |  | 1.1572 | 0.49396 | 1.4006 | 0.67157 | 0.79352 | 0.31481 | 0.33714 | 0.69761 | 0.45898 | 0.30948 | 0.32967 | 0.20976 | 1.2712
| Rel.Std.Dev. |  | 3.439% | 1.52% | 4.146% | 2.014% | 2.377% | 0.9217% | 1.021% | 2.091% | 1.325% | 0.9213% | 0.9185% | 0.6337% | 3.852%
| PDM3 |  | 0.4311% | -3.001% | 0.8291% | -0.4642% | -0.3648% | 1.938% | -1.459% | -0.4145% | 3.366% | 0.257% | 7.121% | -1.21% | -1.509%
| Z-Score (Absolute) |  | 0.49 | 1.62 | 1.72 | 0.06 | 0 | 1.42 | 0.67 | 0.03 | 2.3 | 0.38 | 4.61 | 0.52 | 0.7
| Analytical results for Er in OREAS 461 (Certified Value 7.46 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 7.4 | 7.4 | 9.14 | 7.33 | 7.31 | 7.2592 | 7.3 | 7.3 | 11.1 | 7.14 | 12.5661 | 7.4 | 8.45 |  | 1
| 1 | 2 | 8.1 | 7.44 | 9.05 | 7.42 | 7.08 | 7.6145 | 7.2 | 7.3 | 11.4 | 7.37 | 12.4829 | 7 | 7.75 |  | 6
| 1 | 3 | 7.97 | 7.41 | 9.16 | 7.45 | 7.4 | 7.6086 | 7.5 | 7.4 | 11.2 | 7.57 | 12.7797 | 6.95 | 7.8 |  | 16
| 1 | 4 | 8.08 | 7.57 | 8.47 | 7.42 | 7.07 | 7.4815 | 7.6 | 7.4 | 11 | 7.43 | 12.5747 | 7.35 | 7.65 |  | 7.4597
| 1 | 5 | 7.66 | 7.52 | 9.23 | 7.54 | 7.14 | 7.2285 | 7.3 | 6.9 | 11.6 | 7.32 | 12.5352 | 7.35 | 8.65 |  | 82
| 1 | 6 | 7.7 | 7.63 | 8.99 | 7.72 | 7.38 | 7.3954 | 7.4 | 7.2 | 11.2 | 7.54 | 12.7716 | 7.4 | 8.2
| Mean |  | 7.8183 | 7.495 | 9.0067 | 7.48 | 7.23 | 7.4313 | 7.3833 | 7.25 | 11.25 | 7.395 | 12.6184 | 7.2417 | 8.0833
| Median |  | 7.835 | 7.48 | 9.095 | 7.435 | 7.225 | 7.4385 | 7.35 | 7.3 | 11.2 | 7.4 | 12.5704 | 7.35 | 8
| Std Dev. |  | 0.27745 | 0.093541 | 0.27616 | 0.1355 | 0.151 | 0.16709 | 0.1472 | 0.18708 | 0.21679 | 0.15758 | 0.12602 | 0.20837 | 0.41191
| Rel.Std.Dev. |  | 3.549% | 1.248% | 3.066% | 1.811% | 2.088% | 2.248% | 1.994% | 2.58% | 1.927% | 2.131% | 0.9987% | 2.877% | 5.096%
| PDM3 |  | 4.808% | 0.4737% | 20.74% | 0.2726% | -3.079% | -0.3802% | -1.023% | -2.811% | 50.81% | -0.8668% | 69.15% | -2.922% | 8.361%
| Z-Score (Absolute) |  | 1.37 | 0.22 | 5.96 | 0 | 0.72 | 0 | 0.17 | 0.4 | 13.53 | 0.13 | 18.37 | 0.67 | 1.48
| Analytical results for Eu in OREAS 461 (Certified Value 47.6 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 51.2 | 46.1 | 50.4 | 48.7 | 47.6 | 47.6856 | 46 | 46.9 | 44.2 | 44.11 | 53.2676 | 45.2 | 48.5 |  | 1
| 1 | 2 | 55.1 | 46.4 | 48.6 | 49 | 46.5 | 47.7181 | 46.1 | 48.6 | 45.2 | 43.9 | 53.3125 | 45.4 | 39 |  | 7
| 1 | 3 | 55 | 47.1 | 49.3 | 50 | 47.8 | 47.846 | 46.9 | 48.2 | 44.7 | 43.76 | 53.2948 | 45.8 | 44 |  | 16
| 1 | 4 | 55.8 | 45.6 | 45.2 | 49.7 | 46.5 | 47.6537 | 47.3 | 47.4 | 43.2 | 44.4 | 53.1393 | 45.1 | 43.4 |  | 47.6394
| 1 | 5 | 52.7 | 46 | 49.5 | 49.8 | 47.8 | 47.8968 | 46.1 | 48.3 | 43 | 43.94 | 53.127 | 46 | 44 |  | 83
| 1 | 6 | 51.4 | 44.8 | 48.2 | 51.2 | 48.4 | 47.9085 | 46.6 | 49.1 | 44.3 | 42.56 | 53.5906 | 46.3 | 45.1
| Mean |  | 53.5333 | 46 | 48.5333 | 49.7333 | 47.4333 | 47.7848 | 46.5 | 48.0833 | 44.1 | 43.7783 | 53.2886 | 45.6333 | 44
| Median |  | 53.85 | 46.05 | 48.95 | 49.75 | 47.7 | 47.782 | 46.35 | 48.25 | 44.25 | 43.92 | 53.2812 | 45.6 | 44
| Std Dev. |  | 2.0216 | 0.77201 | 1.8019 | 0.8756 | 0.77115 | 0.11232 | 0.52536 | 0.80353 | 0.85323 | 0.63594 | 0.16773 | 0.4761 | 3.0601
| Rel.Std.Dev. |  | 3.776% | 1.678% | 3.713% | 1.761% | 1.626% | 0.2351% | 1.13% | 1.671% | 1.935% | 1.453% | 0.3148% | 1.043% | 6.955%
| PDM3 |  | 12.37% | -3.441% | 1.876% | 4.395% | -0.4325% | 0.3052% | -2.392% | 0.9319% | -7.43% | -8.105% | 11.86% | -4.211% | -7.639%
| Z-Score (Absolute) |  | 2.29 | 0.54 | 0.66 | 0.86 | 0 | 0.13 | 0.35 | 0.24 | 1.25 | 1.28 | 2.19 | 0.67 | 1.29
| Analytical results for Fe in OREAS 461 (Certified Value 31.52 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 30 | 30.7 | 30.3 | 33.1 | 32.4 | 32.426 | 30.51 | 31.2 | 31.9 | 31.07 | 31.2398 | > 15 | 32.0185 | 31.9 | 1
| 1 | 2 | 29.9 | 30.5 | 30.2 | 32.6 | 32.2 | 32.4677 | 30.22 | 30.95 | 32.1 | 31.48 | 31.5378 | > 15 | 31.892 | 31.9 | #N/A
| 1 | 3 | 27.7 | 30.8 | 30.2 | 33.3 | 30.3 | 32.7117 | 30.08 | 31.52 | 32.1 | 31.14 | 31.0354 | > 15 | 31.825 | 32.3 | 16
| 1 | 4 | 28.2 | 30.8 | 30.3 | 32.9 | 29.7 | 32.8824 | 30.49 | 31.3 | 32 | 32.42 | 31.5479 | > 15 | 31.539 | 31.5 | 31.5178
| 1 | 5 | 27.7 | 30.8 | 30.1 | 33.7 | 31.5 | 32.9889 | 30.3 | 31.41 | 32.1 | 32.24 | 31.3844 | > 15 | 31.59 | 31.5 | 84
| 1 | 6 | 28.2 | 31 | 30.2 | 33.5 | 29.9 | 32.9064 | 30.4 | 31.76 | 32 | 30.96 | 30.9791 | > 15 | 31.76 | 32.8
| Mean |  | 28.6167 | 30.7667 | 30.2167 | 33.1833 | 31 | 32.7305 | 30.3333 | 31.3567 | 32.0333 | 31.5517 | 31.2874 |  | 31.7707 | 31.9833
| Median |  | 28.2 | 30.8 | 30.2 | 33.2 | 30.9 | 32.797 | 30.35 | 31.355 | 32.05 | 31.31 | 31.3121 |  | 31.7925 | 31.9
| Std Dev. |  | 1.0572 | 0.1633 | 0.075277 | 0.40208 | 1.1866 | 0.23788 | 0.16633 | 0.27761 | 0.08165 | 0.63 | 0.24527 |  | 0.18194 | 0.49967
| Rel.Std.Dev. |  | 3.694% | 0.5308% | 0.2491% | 1.212% | 3.828% | 0.7268% | 0.5484% | 0.8853% | 0.2549% | 1.997% | 0.7839% |  | 0.5727% | 1.562%
| PDM3 |  | -9.205% | -2.383% | -4.128% | 5.284% | -1.643% | 3.848% | -3.758% | -0.5113% | 1.636% | 0.1074% | -0.731% |  | 0.8025% | 1.477%
| Z-Score (Absolute) |  | 2.95 | 0.63 | 1.23 | 1.97 | 0.38 | 1.48 | 1.1 | 0 | 0.73 | 0.21 | 0.07 | NA | 0.45 | 0.67
| Analytical results for Ga in OREAS 461 (Certified Value 45.5 ppm) | [15 empty] | Certified
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 10; 08: 12; 09: 22 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS | 1
1
| 1 | 1 | 01: 47.8; 02: 47.4; 03: 47.9; 05: 60.8; 06: 35.7685; 07: 10.7; 08: 46.3011; 09: 35.4 | 1
| 1 | 2 | 01: 47.1; 02: 47.9; 03: 47.3; 05: 58.2; 06: 37.1335; 07: 10.5; 08: 46.1081; 09: 33.4 | #N/A
| 1 | 3 | 01: 48.2; 02: 48.4; 03: 47.7; 05: 60; 06: 35.9095; 07: 10.1; 08: 46.3029; 09: 34.6 | 16
| 1 | 4 | 01: 48.4; 02: 43.8; 03: 47.3; 05: 57.9; 06: 36; 07: 10.6; 08: 45.8842; 09: 33.8 | 45.5202
| 1 | 5 | 01: 47.2; 02: 48.3; 03: 48.9; 05: 59.3; 06: 34.9198; 07: 10.2; 08: 45.6738; 09: 34 | 85
| 1 | 6 | 01: 47; 02: 46.1; 03: 48.6; 05: 59.5; 06: 36.084; 07: 10.7; 08: 46.545; 09: 33.2
| Mean |  | 01: 47.6167; 02: 46.9833; 03: 47.95; 05: 59.2833; 06: 35.9692; 07: 10.4667; 08: 46.1359; 09: 34.0667
| Median |  | 01: 47.5; 02: 47.65; 03: 47.8; 05: 59.4; 06: 35.9548; 07: 10.55; 08: 46.2046; 09: 33.9
| Std Dev. |  | 01: 0.60139; 02: 1.7702; 03: 0.66858; 05: 1.0907; 06: 0.70919; 07: 0.2582; 08: 0.31631; 09: 0.8165
| Rel.Std.Dev. |  | 01: 1.263%; 02: 3.768%; 03: 1.394%; 05: 1.84%; 06: 1.972%; 07: 2.467%; 08: 0.6856%; 09: 2.397%
| PDM3 |  | 01: 4.605%; 02: 3.214%; 03: 5.338%; 05: 30.24%; 06: -20.98%; 07: -77.01%; 08: 1.352%; 09: -25.16%
| Z-Score (Absolute) |  | 01: 0.08; 02: 0.08; 03: 0.12; 05: 1.4; 06: 1.23; 07: 4.1; 08: 0.08; 09: 1.44
| Analytical results for Gd in OREAS 461 (Certified Value 98 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 88.1 | 96.9 | 127.5 | 90.9 | 102.5 | 99.1766 | 99.2 | 96.1 | 47.9 | 96.67 | 142.1213 | 96.4 | 107 |  | 1
| 1 | 2 | 97.5 | 95.3 | 117.5 | 90.7 | 97.3 | 97.0894 | 100.6 | 99.3 | 45.8 | 95.38 | 141.6797 | 95.3 | 98 |  | 8
| 1 | 3 | 96.9 | 95.8 | 120.5 | 93.2 | 103 | 98.2242 | 100.1 | 97.4 | 45.6 | 95.31 | 141.0163 | 97.9 | 104 |  | 16
| 1 | 4 | 100.5 | 103 | 103 | 91.8 | 97.8 | 98.8098 | 103.1 | 96.4 | 46.4 | 96.82 | 141.1854 | 97.1 | 100 |  | 97.5977
| 1 | 5 | 94.7 | 98.1 | 120.5 | 92.2 | 102 | 99.8001 | 99.8 | 100.1 | 44.8 | 94.42 | 141.7589 | 96.8 | 101 |  | 86
| 1 | 6 | 93 | 97.6 | 117 | 94.5 | 102 | 96.8614 | 99.9 | 100.4 | 46.5 | 94.04 | 142.0565 | 98.5 | 103
| Mean |  | 95.1167 | 97.7833 | 117.6667 | 92.2167 | 100.7667 | 98.3269 | 100.45 | 98.2833 | 46.1667 | 95.44 | 141.6363 | 97 | 102.1667
| Median |  | 95.8 | 97.25 | 119 | 92 | 102 | 98.517 | 100 | 98.35 | 46.1 | 95.345 | 141.7193 | 96.95 | 102
| Std Dev. |  | 4.2832 | 2.7651 | 8.1035 | 1.4414 | 2.524 | 1.1675 | 1.3751 | 1.8925 | 1.0482 | 1.1347 | 0.45089 | 1.1278 | 3.1885
| Rel.Std.Dev. |  | 4.503% | 2.828% | 6.887% | 1.563% | 2.505% | 1.187% | 1.369% | 1.926% | 2.27% | 1.189% | 0.3183% | 1.163% | 3.121%
| PDM3 |  | -2.542% | 0.1902% | 20.56% | -5.513% | 3.247% | 0.7472% | 2.923% | 0.7025% | -52.7% | -2.211% | 45.12% | -0.6124% | 4.681%
| Z-Score (Absolute) |  | 0.75 | 0.37 | 5.29 | 1.44 | 0.59 | 0.01 | 0.39 | 0 | 12.36 | 0.67 | 10.28 | 0.3 | 0.92
| Analytical results for Ge in OREAS 461 (Indicative Value 3.6 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 2.35 | 4.11 | 3.61 | 2.73 | 5.3 | 0.5 | 3.8 | 6.7 | < 0.05 | 4.5 |  |  |  |  | 1
| 1 | 2 | 2.51 | 3.93 | 3.47 | 2.63 | 4.79 | 0.43227 | 3.6 | 6.3 | < 0.05 | 4.6 |  |  |  |  | 20
| 1 | 3 | 2.48 | 3.86 | 3.56 | 2.67 | 4.78 | 0.4 | 3.8 | 6.3 | < 0.05 | 4.39 |  |  |  |  | 16
| 1 | 4 | 2.83 | 3.88 | 3.13 | 2.59 | 4.99 | 0.44964 | 3.7 | 5.6 | < 0.05 | 4.42 |  |  |  |  | 3.5955
| 1 | 5 | 2.53 | 3.62 | 3.6 | 2.65 | 4.87 | 0.44734 | 3.7 | 6 | < 0.05 | 4.28 |  |  |  |  | 26
| 1 | 6 | 2.39 | 3.61 | 3.37 | 2.72 | 5.08 | 0.50329 | 4 | 7.1 | < 0.05 | 4.37
| Mean |  | 2.515 | 3.835 | 3.4567 | 2.665 | 4.9683 | 0.45542 | 3.7667 | 6.3333 |  | 4.4267
| Median |  | 2.495 | 3.87 | 3.515 | 2.66 | 4.93 | 0.44849 | 3.75 | 6.3 |  | 4.405
| Std Dev. |  | 0.16944 | 0.19191 | 0.18392 | 0.053572 | 0.19994 | 0.039961 | 0.13663 | 0.52409 |  | 0.11094
| Rel.Std.Dev. |  | 6.737% | 5.004% | 5.321% | 2.01% | 4.024% | 8.774% | 3.627% | 8.275% |  | 2.506%
| PDM3 |  | -30.05% | 6.662% | -3.86% | -25.88% | 38.18% | -87.33% | 4.762% | 76.15% |  | 23.12%
| Z-Score (Absolute) |  | 0.8 | 0.04 | 0.19 | 0.67 | 0.74 | 2.03 | 0 | 1.57 | NA | 0.4
| Analytical results for Hf in OREAS 461 (Certified Value 7.39 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 8.2 | 6.6 | 8.4 | 7.6 | 7.6 | 6.4285 | 6.2 | 7.13 | 8.92 | 6.6 |  |  |  |  | 1
| 1 | 2 | 8.7 | 6.9 | 8 | 7.3 | 7.5 | 6.2507 | 6 | 7.07 | 8.81 | 6.6 |  |  |  |  | 9
| 1 | 3 | 8.6 | 6.7 | 8.3 | 7.6 | 7.6 | 6.7319 | 6.2 | 7.11 | 8.85 | 7 |  |  |  |  | 16
| 1 | 4 | 8.3 | 7.2 | 7.7 | 7.7 | 7.2 | 6.9212 | 5.9 | 6.91 | 9.06 | 6.8 |  |  |  |  | 7.3912
| 1 | 5 | 8.2 | 6.9 | 8.3 | 7.9 | 7.6 | 6.8714 | 6.1 | 6.99 | 8.68 | 6.8 |  |  |  |  | 87
| 1 | 6 | 8.1 | 6.7 | 8 | 7.9 | 7.9 | 6.6261 | 6.2 | 7.25 | 8.46 | 6.8
| Mean |  | 8.35 | 6.8333 | 8.1167 | 7.6667 | 7.5667 | 6.6383 | 6.1 | 7.0767 | 8.7967 | 6.7667
| Median |  | 8.25 | 6.8 | 8.15 | 7.65 | 7.6 | 6.679 | 6.15 | 7.09 | 8.83 | 6.8
| Std Dev. |  | 0.2429 | 0.21602 | 0.26394 | 0.22509 | 0.22509 | 0.25983 | 0.12649 | 0.11776 | 0.20714 | 0.15055
| Rel.Std.Dev. |  | 2.909% | 3.161% | 3.252% | 2.936% | 2.975% | 3.914% | 2.074% | 1.664% | 2.355% | 2.225%
| PDM3 |  | 12.97% | -7.547% | 9.816% | 3.727% | 2.375% | -10.19% | -17.47% | -4.255% | 19.02% | -8.449%
| Z-Score (Absolute) |  | 1.12 | 0.53 | 0.87 | 0.38 | 0.27 | 0.74 | 1.33 | 0.27 | 1.61 | 0.6
| Analytical results for Ho in OREAS 461 (Certified Value 4.06 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.33 | 3.95 | 4.63 | 4.07 | 3.96 | 4.0917 | 4.2 | 3.9 | 5.25 | 4.02 | 5.0412 | 4.1 | 4.26 |  | 1
| 1 | 2 | 4.56 | 4.05 | 4.53 | 4.01 | 3.8 | 4.0883 | 4.2 | 4.2 | 5.33 | 4.04 | 4.9453 | 4 | 3.88 |  | 10
| 1 | 3 | 4.51 | 4.03 | 4.55 | 4.08 | 3.98 | 4.0844 | 4.3 | 4 | 5.19 | 4.05 | 5.0033 | 4.15 | 4 |  | 16
| 1 | 4 | 4.67 | 4.02 | 4.14 | 4.15 | 3.79 | 4.2018 | 4.2 | 3.8 | 5.04 | 4.05 | 4.991 | 4.05 | 3.94 |  | 4.0576
| 1 | 5 | 4.41 | 3.98 | 4.61 | 4.14 | 3.89 | 4.2227 | 4.1 | 4.1 | 5.01 | 4.07 | 5.0459 | 4.1 | 3.94 |  | 88
| 1 | 6 | 4.39 | 4.05 | 4.45 | 4.21 | 3.95 | 4.122 | 4.2 | 4.1 | 5.23 | 4.02 | 4.9731 | 4.1 | 4.12
| Mean |  | 4.4783 | 4.0133 | 4.485 | 4.11 | 3.895 | 4.1352 | 4.2 | 4.0167 | 5.175 | 4.0417 | 5 | 4.0833 | 4.0233
| Median |  | 4.46 | 4.025 | 4.54 | 4.11 | 3.92 | 4.1068 | 4.2 | 4.05 | 5.21 | 4.045 | 4.9971 | 4.1 | 3.97
| Std Dev. |  | 0.12561 | 0.040332 | 0.18064 | 0.070711 | 0.083126 | 0.06152 | 0.063246 | 0.1472 | 0.12518 | 0.019408 | 0.039024 | 0.05164 | 0.14166
| Rel.Std.Dev. |  | 2.805% | 1.005% | 4.028% | 1.72% | 2.134% | 1.488% | 1.506% | 3.665% | 2.419% | 0.4802% | 0.7805% | 1.265% | 3.521%
| PDM3 |  | 10.37% | -1.091% | 10.53% | 1.291% | -4.008% | 1.911% | 3.509% | -1.009% | 27.54% | -0.3929% | 23.22% | 0.6339% | -0.8448%
| Z-Score (Absolute) |  | 2.66 | 0.7 | 3.21 | 0 | 1.55 | 0.18 | 0.65 | 0.67 | 7.69 | 0.49 | 6.43 | 0.19 | 0.63
| Analytical results for In in OREAS 461 (Certified Value 0.57 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 0.557 | 0.568 | 0.615 | 0.561 | 0.586 | 0.5264 | 0.6 | 0.6 | 0.56 | 0.56 | 0.6 | 0.48 |  |  | 1
| 1 | 2 | 0.555 | 0.548 | 0.668 | 0.55 | 0.552 | 0.54801 | 0.6 | 0.58 | 0.557 | 0.54 | 0.6 | 0.52 |  |  | #N/A
| 1 | 3 | 0.555 | 0.581 | 0.693 | 0.575 | 0.574 | 0.519 | 0.6 | 0.5 | 0.545 | 0.57 | 0.6 | 0.5 |  |  | 16
| 1 | 4 | 0.567 | 0.581 | 0.656 | 0.571 | 0.576 | 0.54597 | 0.6 | 0.63 | 0.549 | 0.55 | 0.6 | 0.46 |  |  | 0.56812
| 1 | 5 | 0.565 | 0.569 | 0.676 | 0.585 | 0.588 | 0.5306 | 0.6 | 0.5 | 0.564 | 0.55 | 0.6 | 0.48 |  |  | 89
| 1 | 6 | 0.548 | 0.566 | 0.644 | 0.567 | 0.592 | 0.5343 | 0.6 | 0.6 | 0.561 | 0.53 | 0.6 | 0.48
| Mean |  | 0.55783 | 0.56883 | 0.65867 | 0.56817 | 0.578 | 0.53405 | 0.6 | 0.56833 | 0.556 | 0.55 | 0.6 | 0.48667
| Median |  | 0.556 | 0.5685 | 0.662 | 0.569 | 0.581 | 0.53245 | 0.6 | 0.59 | 0.5585 | 0.55 | 0.6 | 0.48
| Std Dev. |  | 0.0070545 | 0.012156 | 0.027186 | 0.012007 | 0.014533 | 0.011259 | 0 | 0.055287 | 0.0074297 | 0.014142 | 0 | 0.020656
| Rel.Std.Dev. |  | 1.265% | 2.137% | 4.127% | 2.113% | 2.514% | 2.108% | 0% | 9.728% | 1.336% | 2.571% | 0% | 4.244%
| PDM3 |  | -1.811% | 0.1253% | 15.94% | 0.007975% | 1.739% | -5.998% | 5.611% | 0.03731% | -2.134% | -3.19% | 5.611% | -14.34%
| Z-Score (Absolute) |  | 0.46 | 0.03 | 4 | 0 | 0.43 | 1.51 | 1.4 | 0 | 0.54 | 0.81 | 1.4 | 3.9
| Analytical results for K in OREAS 461 (Certified Value 0.233 wt.%) | [15 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.2 | 0.23 | 0.23 | 0.24 | 0.25 | 0.24 | 0.23 | 0.249 | 0.2331 | 0.23 | 0.23 | 0.23 |  |  | 1
| 1 | 2 | 0.22 | 0.29 | 0.23 | 0.23 | 0.24 | 0.23 | 0.23 | 0.245 | 0.2383 | 0.23 | 0.23 | 0.22 |  |  | 11
| 1 | 3 | 0.21 | 0.23 | 0.22 | 0.23 | 0.23 | 0.23 | 0.23 | 0.245 | 0.2359 | 0.24 | 0.23 | 0.22 |  |  | 16
| 1 | 4 | 0.21 | 0.24 | 0.23 | 0.24 | 0.23 | 0.24 | 0.22 | 0.24 | 0.2417 | 0.23 | 0.23 | 0.21 |  |  | 0.23312
| 1 | 5 | 0.2 | 0.23 | 0.23 | 0.24 | 0.24 | 0.24 | 0.24 | 0.245 | 0.2334 | 0.23 | 0.23 | 0.21 |  |  | 90
| 1 | 6 | 0.19 | 0.23 | 0.23 | 0.24 | 0.22 | 0.24 | 0.24 | 0.247 | 0.2343 | 0.23 | 0.23 | 0.22
| Mean |  | 0.205 | 0.24167 | 0.22833 | 0.23667 | 0.235 | 0.23667 | 0.23167 | 0.24517 | 0.23612 | 0.23167 | 0.23 | 0.21833
| Median |  | 0.205 | 0.23 | 0.23 | 0.24 | 0.235 | 0.24 | 0.23 | 0.245 | 0.2351 | 0.23 | 0.23 | 0.22
| Std Dev. |  | 0.010488 | 0.024014 | 0.0040825 | 0.005164 | 0.010488 | 0.005164 | 0.0075277 | 0.0029944 | 0.0033397 | 0.0040825 | 0 | 0.0075277
| Rel.Std.Dev. |  | 5.116% | 9.937% | 1.788% | 2.182% | 4.463% | 2.182% | 3.249% | 1.221% | 1.414% | 1.762% | 0% | 3.448%
| PDM3 |  | -12.06% | 3.664% | -2.055% | 1.52% | 0.8047% | 1.52% | -0.6252% | 5.166% | 1.284% | -0.6252% | -1.34% | -6.345%
| Z-Score (Absolute) |  | 4.65 | 0.03 | 0.61 | 0.84 | 0.55 | 0.84 | 0.03 | 2.31 | 0.74 | 0.03 | 0.32 | 2.34
| Analytical results for La in OREAS 461 (Certified Value 2476 ppm) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 2060 | 2620 | 2130 | 2200 | 2120 | 2680.0044 | 2514.3 | 2509 | 2720 | 2510.28 | 2646.4168 | 2696 | 2530 |  | 1
| 1 | 2 | 2350 | 2600 | 2140 | 2200 | 2190 | 2735.4177 | 2508.4 | 2570 | 2760 | 2529.21 | 2583.5317 | 2656 | 2530 |  | #N/A
| 1 | 3 | 2270 | 2630 | 2120 | 2220 | 2100 | 2692.6341 | 2550.9 | 2519.4 | 2720 | 2531.5 | 2696.8915 | 2702 | 2580 |  | 16
| 1 | 4 | 2330 | 2630 | 2140 | 2220 | 2010 | 2787.476 | 2539.8 | 2521.6 | 2670 | 2581.26 | 2594.2433 | 2703.5 | 2520 |  | 2476.481
| 1 | 5 | 2150 | 2640 | 2120 | 2280 | 2140 | 2801.3478 | 2527.4 | 2548.3 | 2670 | 2517.51 | 2645.7666 | 2701.5 | 2540 |  | 91
| 1 | 6 | 2080 | 2640 | 2130 | 2240 | 2010 | 2696.024 | 2508.9 | 2607.4 | 2690 | 2449.6 | 2679.0053 | 2743 | 2620
| Mean |  | 2206.6667 | 2626.6667 | 2130 | 2226.6667 | 2095 | 2732.1507 | 2524.95 | 2545.95 | 2705 | 2519.8933 | 2640.9759 | 2700.3333 | 2553.3333
| Median |  | 2210 | 2630 | 2130 | 2220 | 2110 | 2715.7209 | 2520.85 | 2534.95 | 2705 | 2523.36 | 2646.0917 | 2701.75 | 2535
| Std Dev. |  | 126.912 | 15.0555 | 8.9443 | 30.1109 | 72.3187 | 51.857 | 17.5776 | 37.4784 | 35.0714 | 42.4891 | 44.9647 | 27.6363 | 38.8158
| Rel.Std.Dev. |  | 5.751% | 0.5732% | 0.4199% | 1.352% | 3.452% | 1.898% | 0.6962% | 1.472% | 1.297% | 1.686% | 1.703% | 1.023% | 1.52%
| PDM3 |  | -10.9% | 6.064% | -13.99% | -10.09% | -15.4% | 10.32% | 1.957% | 2.805% | 9.228% | 1.753% | 6.642% | 9.039% | 3.103%
| Z-Score (Absolute) |  | 1.4 | 0.36 | 1.72 | 1.32 | 1.87 | 0.81 | 0.06 | 0.03 | 0.69 | 0.08 | 0.42 | 0.67 | 0
| Analytical results for Li in OREAS 461 (Certified Value 12 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 1
1
| 1 | 1 | 9.2 | 11.9 | 10.4 | 12.6 | 11.2 | 12.9489 | 12.1 | 13 | 11.8 | 10.9 | 14.2544 | 13 | 10 | 12 | 1
| 1 | 2 | 9.9 | 12.5 | 10.5 | 12.3 | 10.6 | 12.6566 | 12.1 | 13 | 11.8 | 10.4 | 14.187 | 12 | 15 | 11.5 | #N/A
| 1 | 3 | 10 | 12.2 | 11.1 | 12.6 | 11 | 13.3867 | 12.1 | 13 | 12 | 10 | 13.8567 | 13 | 10 | 12 | 16
| 1 | 4 | 10.1 | 11.8 | 12.3 | 12.7 | 10 | 13.0297 | 12.1 | 13 | 11.9 | 10.5 | 13.3227 | 14 | 15 | 11.5 | 12.0397
| 1 | 5 | 10 | 11.7 | 10.9 | 13 | 10.4 | 13.6092 | 11.8 | 14 | 11.6 | 10.5 | 13.5922 | 13 | 15 | 12 | 92
| 1 | 6 | 9.9 | 11.8 | 10.8 | 13 | 10 | 13.2865 | 11.9 | 13 | 12 | 10.4 | 14.1885 | 13 | 15 | 11.5
| Mean |  | 9.85 | 11.9833 | 11 | 12.7 | 10.5333 | 13.1529 | 12.0167 | 13.1667 | 11.85 | 10.45 | 13.9002 | 13 | 13.3333 | 11.75
| Median |  | 9.95 | 11.85 | 10.85 | 12.65 | 10.5 | 13.1581 | 12.1 | 13 | 11.85 | 10.45 | 14.0219 | 13 | 15 | 11.75
| Std Dev. |  | 0.32711 | 0.30605 | 0.68702 | 0.26833 | 0.50067 | 0.3417 | 0.13292 | 0.40825 | 0.15166 | 0.2881 | 0.37978 | 0.63246 | 2.582 | 0.27386
| Rel.Std.Dev. |  | 3.321% | 2.554% | 6.246% | 2.113% | 4.753% | 2.598% | 1.106% | 3.101% | 1.28% | 2.757% | 2.732% | 4.865% | 19.36% | 2.331%
| PDM3 |  | -18.19% | -0.4686% | -8.636% | 5.484% | -12.51% | 9.246% | -0.1917% | 9.36% | -1.576% | -13.2% | 15.45% | 7.976% | 10.74% | -2.407%
| Z-Score (Absolute) |  | 1.17 | 0.01 | 0.73 | 0.41 | 0.85 | 0.67 | 0.01 | 0.68 | 0.09 | 0.9 | 1.1 | 0.58 | 0.78 | 0.15
| Analytical results for Lu in OREAS 461 (Certified Value 0.39 ppm) | [15 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 0.52 | 0.4 | 0.41 | 0.34 | 0.39 | 0.36128 | 0.3 | 0.38 | 0.452 | 0.34 | 0.44422 | 0.3 | 0.42 |  | 1
| 1 | 2 | 0.56 | 0.39 | 0.41 | 0.33 | 0.38 | 0.352 | 0.3 | 0.29 | 0.459 | 0.33 | 0.37692 | 0.3 | 0.4 |  | 1
| 1 | 3 | 0.55 | 0.39 | 0.42 | 0.35 | 0.4 | 0.35677 | 0.3 | 0.3 | 0.466 | 0.34 | 0.45734 | 0.3 | 0.42 |  | 16
| 1 | 4 | 0.54 | 0.42 | 0.39 | 0.36 | 0.39 | 0.36321 | 0.3 | 0.29 | 0.454 | 0.33 | 0.41741 | 0.4 | 0.42 |  | 0.38585
| 1 | 5 | 0.51 | 0.4 | 0.43 | 0.35 | 0.4 | 0.36308 | 0.3 | 0.28 | 0.466 | 0.38 | 0.40154 | 0.3 | 0.42 |  | 93
| 1 | 6 | 0.51 | 0.39 | 0.4 | 0.37 | 0.4 | 0.37721 | 0.4 | 0.32 | 0.473 | 0.33 | 0.40907 | 0.4 | 0.44
| Mean |  | 0.53167 | 0.39833 | 0.41 | 0.35 | 0.39333 | 0.36226 | 0.31667 | 0.31 | 0.46167 | 0.34167 | 0.41775 | 0.33333 | 0.42
| Median |  | 0.53 | 0.395 | 0.41 | 0.35 | 0.395 | 0.36218 | 0.3 | 0.295 | 0.4625 | 0.335 | 0.41324 | 0.3 | 0.42
| Std Dev. |  | 0.02137 | 0.01169 | 0.014142 | 0.014142 | 0.008165 | 0.008502 | 0.040825 | 0.036878 | 0.0080664 | 0.019408 | 0.029235 | 0.05164 | 0.012649
| Rel.Std.Dev. |  | 4.019% | 2.935% | 3.449% | 4.041% | 2.076% | 2.347% | 12.89% | 11.9% | 1.747% | 5.68% | 6.998% | 15.49% | 3.012%
| PDM3 |  | 37.79% | 3.236% | 6.26% | -9.29% | 1.94% | -6.113% | -17.93% | -19.66% | 19.65% | -11.45% | 8.268% | -13.61% | 8.851%
| Z-Score (Absolute) |  | 2.15 | 0.08 | 0.26 | 0.67 | 0 | 0.48 | 1.45 | 1.51 | 1.06 | 0.8 | 0.38 | 0.93 | 0.41
| Analytical results for Mg in OREAS 461 (Certified Value 0.999 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.89 | 0.94 | 0.83 | 0.88 | 0.83 | 1.1421 | 1.01 | 1.04 | 1.08 | 1.1136 | 1.069 | 0.91 | 1.082 | 1.09 | 1
| 1 | 2 | 0.97 | 0.93 | 0.83 | 0.87 | 0.85 | 1.1798 | 1.01 | 1.04 | 1.09 | 1.1414 | 1.0852 | 0.92 | 1.0815 | 1.1 | 2
| 1 | 3 | 0.93 | 0.95 | 0.83 | 0.88 | 0.82 | 1.1297 | 0.99 | 1.05 | 1.09 | 1.1226 | 1.0593 | 0.92 | 1.0785 | 1.12 | 16
| 1 | 4 | 0.95 | 0.95 | 0.83 | 0.87 | 0.78 | 1.135 | 1.01 | 1.03 | 1.09 | 1.1684 | 1.0515 | 0.94 | 1.075 | 1.1 | 0.99901
| 1 | 5 | 0.9 | 0.95 | 0.83 | 0.91 | 0.84 | 1.1427 | 1.02 | 1.04 | 1.09 | 1.1137 | 1.052 | 0.93 | 1.0825 | 1.11 | 94
| 1 | 6 | 0.89 | 0.95 | 0.83 | 0.9 | 0.78 | 1.1812 | 1.01 | 1.06 | 1.08 | 1.118 | 1.0535 | 0.95 | 1.0755 | 1.12
| Mean |  | 0.92167 | 0.945 | 0.83 | 0.885 | 0.81667 | 1.1518 | 1.0083 | 1.0433 | 1.0867 | 1.1296 | 1.0617 | 0.92833 | 1.0792 | 1.1067
| Median |  | 0.915 | 0.95 | 0.83 | 0.88 | 0.825 | 1.1424 | 1.01 | 1.04 | 1.09 | 1.1203 | 1.0564 | 0.925 | 1.08 | 1.105
| Std Dev. |  | 0.033714 | 0.0083666 | 0 | 0.016432 | 0.030111 | 0.022785 | 0.0098319 | 0.010328 | 0.005164 | 0.02162 | 0.013236 | 0.01472 | 0.0033417 | 0.012111
| Rel.Std.Dev. |  | 3.658% | 0.8854% | 0% | 1.857% | 3.687% | 1.978% | 0.9751% | 0.9899% | 0.4752% | 1.914% | 1.247% | 1.586% | 0.3097% | 1.094%
| PDM3 |  | -7.742% | -5.407% | -16.92% | -11.41% | -18.25% | 15.29% | 0.9329% | 4.436% | 8.774% | 13.07% | 6.279% | -7.075% | 8.023% | 10.78%
| Z-Score (Absolute) |  | 0.79 | 0.62 | 1.49 | 1.07 | 1.59 | 0.96 | 0.13 | 0.13 | 0.46 | 0.73 | 0.27 | 0.74 | 0.41 | 0.62
| Analytical results for Mn in OREAS 461 (Certified Value 547 ppm) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 500 | 543 | 421 | 422 | 435 | 580.4329 | 632.7 | 530 | 548 | 629 | 835 | 580 | 670 |  | 1
| 1 | 2 | 546 | 551 | 421 | 422 | 466 | 579.6071 | 630.8 | 530 | 552 | 650 | 838 | 580 | 664 |  | 3
| 1 | 3 | 515 | 549 | 412 | 441 | 461 | 577.9823 | 634.2 | 547 | 554 | 633 | 843 | 580 | 686 |  | 16
| 1 | 4 | 517 | 572 | 418 | 440 | 434 | 596.4188 | 627.6 | 539 | 553 | 655 | 861 | 575 | 664 |  | 547.2361
| 1 | 5 | 517 | 550 | 423 | 466 | 445 | 599.495 | 617.8 | 543 | 566 | 625 | 820 | 580 | 658 |  | 95
| 1 | 6 | 517 | 546 | 425 | 463 | 412 | 597.4664 | 625.5 | 553 | 562 | 628 | 845 | 590 | 684
| Mean |  | 518.6667 | 551.8333 | 420 | 442.3333 | 442.1667 | 588.5671 | 628.1 | 540.3333 | 555.8333 | 636.6667 | 840.3333 | 580.8333 | 671
| Median |  | 517 | 549.5 | 421 | 440.5 | 440 | 588.4258 | 629.2 | 541 | 553.5 | 631 | 840.5 | 580 | 667
| Std Dev. |  | 14.9488 | 10.3037 | 4.5607 | 19.0858 | 19.7931 | 10.1858 | 5.9793 | 9.2448 | 6.7651 | 12.628 | 13.4412 | 4.916 | 11.5065
| Rel.Std.Dev. |  | 2.882% | 1.867% | 1.086% | 4.315% | 4.476% | 1.731% | 0.952% | 1.711% | 1.217% | 1.983% | 1.6% | 0.8464% | 1.715%
| PDM3 |  | -5.221% | 0.8401% | -23.25% | -19.17% | -19.2% | 7.553% | 14.78% | -1.261% | 1.571% | 16.34% | 53.56% | 6.139% | 22.62%
| Z-Score (Absolute) |  | 0.4 | 0.07 | 1.27 | 1.06 | 1.06 | 0.31 | 0.67 | 0.14 | 0 | 0.75 | 2.65 | 0.23 | 1.07
| Analytical results for Mo in OREAS 461 (Certified Value 48.2 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 1
2
| 1 | 1 | 47.1 | 47.8 | 53.9 | 48.5 | 50 | 44.3617 | 46.2 | 50 | 48.6 | 49.5 | 44.3514 | 42 |  |  | 1
| 1 | 2 | 46.4 | 49.3 | 53.5 | 48.1 | 48.4 | 43.7558 | 46.7 | 49 | 48.5 | 49.9 | 44.4944 | 41.5 |  |  | #N/A
| 1 | 3 | 46.1 | 48.9 | 55.4 | 48.7 | 51.2 | 45.1128 | 47.3 | 48 | 48.8 | 50.2 | 44.8468 | 42 |  |  | 16
| 1 | 4 | 47.2 | 48.5 | 50.5 | 49.3 | 49.5 | 45.9679 | 46.8 | 50 | 49 | 49.8 | 44.8924 | 41.5 |  |  | 48.2464
| 1 | 5 | 46 | 47.6 | 54.8 | 50.1 | 49.4 | 45.3298 | 46.5 | 47 | 47.8 | 49.1 | 44.159 | 42 |  |  | 96
| 1 | 6 | 47.1 | 47.5 | 52.9 | 49.6 | 51.1 | 44.8604 | 46.7 | 49 | 48.8 | 49.6 | 44.9333 | 42
| Mean |  | 46.65 | 48.2667 | 53.5 | 49.05 | 49.9333 | 44.8981 | 46.7 | 48.8333 | 48.5833 | 49.6833 | 44.6129 | 41.8333
| Median |  | 46.75 | 48.15 | 53.7 | 49 | 49.75 | 44.9866 | 46.7 | 49 | 48.7 | 49.7 | 44.6706 | 42
| Std Dev. |  | 0.54681 | 0.74476 | 1.7216 | 0.74766 | 1.0764 | 0.77082 | 0.36332 | 1.169 | 0.42151 | 0.37639 | 0.3237 | 0.2582
| Rel.Std.Dev. |  | 1.172% | 1.543% | 3.218% | 1.524% | 2.156% | 1.717% | 0.778% | 2.394% | 0.8676% | 0.7576% | 0.7256% | 0.6172%
| PDM3 |  | -3.309% | 0.04191% | 10.89% | 1.666% | 3.496% | -6.94% | -3.205% | 1.216% | 0.6983% | 2.978% | -7.531% | -13.29%
| Z-Score (Absolute) |  | 0.74 | 0.07 | 2.12 | 0.26 | 0.63 | 1.47 | 0.72 | 0.17 | 0.07 | 0.52 | 1.59 | 2.75
| Analytical results for Na in OREAS 461 (Certified Value 0.178 wt.%) | [15 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.17 | 0.16 | 0.17 | 0.19 | 0.18 | 0.18 | 0.17 | 0.184 | 0.183 | 0.20266 | 0.17 | 0.17 | 0.2 |  | 1
| 1 | 2 | 0.17 | 0.36 | 0.17 | 0.19 | 0.18 | 0.17 | 0.17 | 0.184 | 0.1859 | 0.20186 | 0.17 | 0.17 | 0.2 |  | 4
| 1 | 3 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.17 | 0.18 | 0.184 | 0.1858 | 0.20449 | 0.17 | 0.17 | 0.2 |  | 16
| 1 | 4 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.17 | 0.181 | 0.1918 | 0.20443 | 0.18 | 0.17 | 0.19 |  | 0.17817
| 1 | 5 | 0.15 | 0.16 | 0.17 | 0.19 | 0.18 | 0.18 | 0.17 | 0.184 | 0.1826 | 0.20465 | 0.17 | 0.17 | 0.19 |  | 97
| 1 | 6 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.17 | 0.186 | 0.1847 | 0.20274 | 0.17 | 0.17 | 0.2
| Mean |  | 0.16167 | 0.19333 | 0.17 | 0.19 | 0.175 | 0.17667 | 0.17167 | 0.18383 | 0.18563 | 0.20347 | 0.17167 | 0.17 | 0.19667
| Median |  | 0.16 | 0.16 | 0.17 | 0.19 | 0.175 | 0.18 | 0.17 | 0.184 | 0.18525 | 0.20358 | 0.17 | 0.17 | 0.2
| Std Dev. |  | 0.0075277 | 0.08165 | 0 | 3.0405e-17 | 0.0054772 | 0.005164 | 0.0040825 | 0.0016021 | 0.0033206 | 0.001195 | 0.0040825 | 3.0405e-17 | 0.005164
| Rel.Std.Dev. |  | 4.656% | 42.23% | 0% | 1.6e-14% | 3.13% | 2.923% | 2.378% | 0.8715% | 1.789% | 0.5873% | 2.378% | 1.789e-14% | 2.626%
| PDM3 |  | -9.265% | 8.508% | -4.588% | 6.637% | -1.782% | -0.8464% | -3.653% | 3.176% | 4.186% | 14.2% | -3.653% | -4.588% | 10.38%
| Z-Score (Absolute) |  | 1.02 | 1.15 | 0.38 | 1.15 | 0 | 0.13 | 0.25 | 0.67 | 0.81 | 2.17 | 0.25 | 0.38 | 1.65
| Analytical results for Nb in OREAS 461 (Indicative Value 1153 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 1219.5947 | 1186.1 | 906 | > 500 | 1290 |  |  |  |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 1209.887 | 1177.1 | 871 | > 500 | 1280 |  |  |  |  | 5
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 1237 | 1184.5 | 930 | > 500 | 1360 |  |  |  |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 1244.61 | 1150.5 | 920 | > 500 | 1320 |  |  |  |  | 1152.7857
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 1268.403 | 1123.2 | 912 | > 500 | 1330 |  |  |  |  | 26
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 1238.4612 | 1140.5 | 908 | > 500 | 1260
| Mean |  | 07: 1236.326; 08: 1160.3167; 09: 907.8333; 22: 1306.6667
| Median |  | 07: 1237.7306; 08: 1163.8; 09: 910; 22: 1305
| Std Dev. |  | 07: 20.4036; 08: 26.0692; 09: 20.084; 22: 36.697
| Rel.Std.Dev. |  | 07: 1.65%; 08: 2.247%; 09: 2.212%; 22: 2.808%
| PDM3 |  | 07: 7.247%; 08: 0.6533%; 09: -21.25%; 22: 13.35%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.35 | 0.35 | 2.68 | NA | 1
| Analytical results for Nd in OREAS 461 (Certified Value 1623 ppm) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 1505 | 1720 | 1520 | 1650 | 1550 | 1749.267 | 1597.8 | 1546.2 | 1760 | 1593.22 | 1523.459 | 1675 | 1710 |  | 1
| 1 | 2 | 1625 | 1645 | 1580 | 1635 | 1495 | 1707.1995 | 1561.3 | 1612.6 | 1770 | 1552.38 | 1494.7687 | 1661 | 1590 |  | #N/A
| 1 | 3 | 1610 | 1710 | 1575 | 1680 | 1550 | 1710.5504 | 1615.7 | 1592.5 | 1730 | 1595.91 | 1563.0645 | 1670.5 | 1610 |  | 16
| 1 | 4 | 1675 | 1720 | 1540 | 1755 | 1505 | 1748.9323 | 1620.2 | 1587.8 | 1720 | 1584.81 | 1485.6337 | 1670.5 | 1570 |  | 1623.483
| 1 | 5 | 1555 | 1680 | 1605 | 1775 | 1525 | 1664.5483 | 1587.2 | 1618 | 1700 | 1555.3 | 1526.8783 | 1655.5 | 1570 |  | 98
| 1 | 6 | 1505 | 1690 | 1570 | 1725 | 1555 | 1741.3026 | 1593.8 | 1641.2 | 1750 | 1504.48 | 1550.6707 | 1692.5 | 1670
| Mean |  | 1579.1667 | 1694.1667 | 1565 | 1703.3333 | 1530 | 1720.3 | 1596 | 1599.7167 | 1738.3333 | 1564.35 | 1524.0791 | 1670.8333 | 1620
| Median |  | 1582.5 | 1700 | 1572.5 | 1702.5 | 1537.5 | 1725.9265 | 1595.8 | 1602.55 | 1740 | 1570.055 | 1525.1686 | 1670.5 | 1600
| Std Dev. |  | 69.0229 | 29.0545 | 30.3315 | 57.1548 | 25.6905 | 33.0906 | 21.276 | 32.4938 | 26.3944 | 34.8072 | 30.2402 | 12.7971 | 57.6194
| Rel.Std.Dev. |  | 4.371% | 1.715% | 1.938% | 3.355% | 1.679% | 1.924% | 1.333% | 2.031% | 1.518% | 2.225% | 1.984% | 0.7659% | 3.557%
| PDM3 |  | -2.73% | 4.354% | -3.602% | 4.918% | -5.758% | 5.964% | -1.693% | -1.464% | 7.074% | -3.642% | -6.123% | 2.917% | -0.2145%
| Z-Score (Absolute) |  | 0.2 | 0.91 | 0.34 | 1 | 0.67 | 1.17 | 0.04 | 0 | 1.34 | 0.34 | 0.73 | 0.69 | 0.2
| Analytical results for Ni in OREAS 461 (Certified Value 71 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
1
| 1 | 1 | 72 | 68.8 | 70.9 | 71.3 | 69.3 | 65.2084 | 72 | 72 | 77 | 70 | 86 |  |  |  | 1
| 1 | 2 | 72 | 67.2 | 72.8 | 68.9 | 67.3 | 67.5234 | 74 | 72.6 | 73 | 70 | 86 |  |  |  | #N/A
| 1 | 3 | 72 | 68.9 | 73.1 | 71.9 | 69.2 | 66.2823 | 77 | 71.3 | 75 | 70 | 94 |  |  |  | 16
| 1 | 4 | 73 | 69.3 | 69 | 69.1 | 65.2 | 69.8188 | 82 | 71.3 | 79 | 70 | 94 |  |  |  | 71.1258
| 1 | 5 | 72.5 | 67.1 | 72.3 | 72.1 | 66.7 | 69.0167 | 88 | 70.9 | 75 | 70 | 86 |  |  |  | 99
| 1 | 6 | 72 | 69 | 72.4 | 72.2 | 66.8 | 65.5968 | 75 | 72.4 | 75 | 70 | 86
| Mean |  | 72.25 | 68.3833 | 71.75 | 70.9167 | 67.4167 | 67.2411 | 78 | 71.75 | 75.6667 | 70 | 88.6667
| Median |  | 72 | 68.85 | 72.35 | 71.6 | 67.05 | 66.9029 | 76 | 71.65 | 75 | 70 | 86
| Std Dev. |  | 0.41833 | 0.9704 | 1.545 | 1.5184 | 1.5842 | 1.8779 | 5.9666 | 0.68337 | 2.0656 | 0 | 4.1312
| Rel.Std.Dev. |  | 0.579% | 1.419% | 2.153% | 2.141% | 2.35% | 2.793% | 7.649% | 0.9524% | 2.73% | 0% | 4.659%
| PDM3 |  | 1.581% | -3.856% | 0.8776% | -0.294% | -5.215% | -5.462% | 9.665% | 0.8776% | 6.384% | -1.583% | 24.66%
| Z-Score (Absolute) |  | 0.1 | 0.7 | 0.11 | 0.17 | 0.9 | 0.94 | 0.88 | 0 | 0.67 | 0.36 | 3.51
| Analytical results for P in OREAS 461 (Certified Value 0.338 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.333 | 0.345 | 0.3 | 0.324 | 0.312 | 0.3411 | 0.37 | 0.337 | 0.3628 | 0.39836 | 0.3274 | 0.3635 | 0.335 |  | 1
| 1 | 2 | 0.336 | 0.344 | 0.301 | 0.322 | 0.321 | 0.3409 | 0.37 | 0.336 | 0.3622 | 0.40359 | 0.3123 | 0.365 | 0.345 |  | 6
| 1 | 3 | 0.322 | 0.348 | 0.301 | 0.329 | 0.307 | 0.3376 | 0.37 | 0.338 | 0.3624 | 0.39544 | 0.3253 | 0.3635 | 0.355 |  | 16
| 1 | 4 | 0.328 | 0.348 | 0.301 | 0.326 | 0.294 | 0.3431 | 0.37 | 0.33 | 0.3722 | 0.39672 | 0.3238 | 0.363 | 0.345 |  | 0.33785
| 1 | 5 | 0.32 | 0.349 | 0.3 | 0.335 | 0.316 | 0.3429 | 0.37 | 0.335 | 0.357 | 0.40116 | 0.3217 | 0.36 | 0.355 |  | 100
| 1 | 6 | 0.324 | 0.35 | 0.3 | 0.332 | 0.294 | 0.3433 | 0.38 | 0.337 | 0.35 | 0.39796 | 0.3234 | 0.364 | 0.345
| Mean |  | 0.32717 | 0.34733 | 0.3005 | 0.328 | 0.30733 | 0.34148 | 0.37167 | 0.3355 | 0.3611 | 0.39887 | 0.32232 | 0.36317 | 0.34667
| Median |  | 0.326 | 0.348 | 0.3005 | 0.3275 | 0.3095 | 0.342 | 0.37 | 0.3365 | 0.3623 | 0.39816 | 0.3236 | 0.3635 | 0.345
| Std Dev. |  | 0.0063377 | 0.0023381 | 0.00054772 | 0.0049396 | 0.011308 | 0.0021674 | 0.0040825 | 0.002881 | 0.0073351 | 0.0029997 | 0.0052693 | 0.0016931 | 0.0075277
| Rel.Std.Dev. |  | 1.937% | 0.6732% | 0.1823% | 1.506% | 3.679% | 0.6347% | 1.098% | 0.8587% | 2.031% | 0.752% | 1.635% | 0.4662% | 2.171%
| PDM3 |  | -3.163% | 2.806% | -11.06% | -2.916% | -9.033% | 1.075% | 10.01% | -0.6965% | 6.881% | 18.06% | -4.599% | 7.492% | 2.609%
| Z-Score (Absolute) |  | 0.56 | 0.23 | 1.61 | 0.53 | 1.34 | 0 | 1.19 | 0.24 | 0.77 | 2.25 | 0.67 | 0.85 | 0.2
| Analytical results for Pb in OREAS 461 (Certified Value 104 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 101.5 | 101 | 113 | 104.5 | 108.5 | 104.4025 | 101 | 108 | 90.9 | 173.189 | 140 | 97 |  |  | 1
| 1 | 2 | 102 | 105 | 113.5 | 104 | 106 | 106.0689 | 102 | 108 | 92.4 | 172.839 | 140 | 100 |  |  | #N/A
| 1 | 3 | 98.4 | 103 | 114 | 109 | 107.5 | 105.7455 | 104 | 108 | 95 | 169.819 | 140 | 103 |  |  | 16
| 1 | 4 | 101 | 107 | 112.5 | 107.5 | 104.5 | 107.7794 | 103 | 105 | 91.5 | 170.853 | 130 | 102 |  |  | 103.9184
| 1 | 5 | 96.7 | 103.5 | 117 | 109 | 106 | 107.1044 | 102 | 107 | 106.7 | 169.541 | 150 | 100 |  |  | 101
| 1 | 6 | 99.6 | 103 | 113 | 108 | 106 | 105.4828 | 104 | 108 | 87.8 | 171.023 | 140 | 106
| Mean |  | 99.8667 | 103.75 | 113.8333 | 107 | 106.4167 | 106.0972 | 102.6667 | 107.3333 | 94.05 | 171.2107 | 140 | 101.3333
| Median |  | 100.3 | 103.25 | 113.25 | 107.75 | 106 | 105.9072 | 102.5 | 108 | 91.95 | 170.938 | 140 | 101
| Std Dev. |  | 2.0373 | 2.0433 | 1.633 | 2.2136 | 1.3934 | 1.2015 | 1.2111 | 1.2111 | 6.6202 | 1.5135 | 6.3246 | 3.0768
| Rel.Std.Dev. |  | 2.04% | 1.969% | 1.435% | 2.069% | 1.309% | 1.132% | 1.18% | 1.128% | 7.039% | 0.884% | 4.518% | 3.036%
| PDM3 |  | -3.899% | -0.162% | 9.541% | 2.965% | 2.404% | 2.097% | -1.205% | 3.286% | -9.496% | 64.75% | 34.72% | -2.488%
| Z-Score (Absolute) |  | 1.01 | 0.4 | 1.1 | 0.12 | 0.03 | 0.03 | 0.57 | 0.17 | 2.33 | 10.29 | 5.34 | 0.78
| Analytical results for Pr in OREAS 461 (Certified Value 482 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 472 | 463 | 464 | 478 | 461 | 507.7527 | 466.4 | 465.68 | 502 | 473.31 | 493.3063 | 493.1 | 496 |  | 1
| 1 | 2 | 507 | 461 | 484 | 474 | 436 | 508.1749 | 465.8 | 482.07 | 507 | 477.84 | 492.056 | 480.8 | 454 |  | #N/A
| 1 | 3 | 507 | 469 | 490 | 483 | 455 | 506.6946 | 477.7 | 473.98 | 507 | 489.77 | 491.9507 | 488.3 | 463 |  | 16
| 1 | 4 | 528 | 467 | 472 | 509 | 439 | 522.013 | 476.8 | 474.55 | 501 | 490.32 | 490.0218 | 489.7 | 454 |  | 482.2815
| 1 | 5 | 488 | 456 | 496 | 515 | 446 | 519.6 | 474.9 | 484.41 | 516 | 475.92 | 493.2285 | 488.7 | 459 |  | 102
| 1 | 6 | 471 | 451 | 482 | 503 | 458 | 505.5371 | 473.1 | 488.12 | 501 | 466.57 | 490.684 | 496.5 | 473
| Mean |  | 495.5 | 461.1667 | 481.3333 | 493.6667 | 449.1667 | 511.6287 | 472.45 | 478.135 | 505.6667 | 478.955 | 491.8746 | 489.5167 | 466.5
| Median |  | 497.5 | 462 | 483 | 493 | 450.5 | 507.9638 | 474 | 478.31 | 504.5 | 476.88 | 492.0033 | 489.2 | 461
| Std Dev. |  | 22.4922 | 6.7651 | 11.7075 | 17.4547 | 10.3811 | 7.208 | 5.1718 | 8.2438 | 5.785 | 9.4004 | 1.3245 | 5.2871 | 16.0842
| Rel.Std.Dev. |  | 4.539% | 1.467% | 2.432% | 3.536% | 2.311% | 1.409% | 1.095% | 1.724% | 1.144% | 1.963% | 0.2693% | 1.08% | 3.448%
| PDM3 |  | 2.741% | -4.378% | -0.1966% | 2.361% | -6.866% | 6.085% | -2.039% | -0.8598% | 4.849% | -0.6897% | 1.989% | 1.5% | -3.272%
| Z-Score (Absolute) |  | 0.77 | 1.1 | 0 | 0.67 | 1.76 | 1.66 | 0.49 | 0.17 | 1.33 | 0.13 | 0.58 | 0.45 | 1.13
| Analytical results for Rb in OREAS 461 (Certified Value 13.3 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 7.2 | 14.2 | 14.3 | 13.1 | 13.6 | 12.7533 | 13.2 | 15.2 | 11.4 | 14.36 | 17.7266 | 13 | 14.4 |  | 1
| 1 | 2 | 12.4 | 13.9 | 14.3 | 12.9 | 12.9 | 12.3943 | 13 | 14.2 | 11.3 | 14.14 | 17.2784 | 13 | 13.6 |  | #N/A
| 1 | 3 | 13.2 | 14.1 | 14.6 | 13 | 13.5 | 12.6891 | 12.9 | 14.7 | 11.5 | 14.43 | 17.2067 | 13 | 13.8 |  | 16
| 1 | 4 | 15.6 | 14.3 | 13.7 | 13.3 | 13 | 12.9409 | 13.1 | 14 | 11.4 | 14.35 | 17.5604 | 13 | 13.6 |  | 13.2859
| 1 | 5 | 9.7 | 13.9 | 14.4 | 13.4 | 13.3 | 12.6223 | 13 | 13.9 | 11.3 | 14 | 17.8282 | 12.5 | 13.6 |  | 103
| 1 | 6 | 8.4 | 14.1 | 13.8 | 13.4 | 13.4 | 12.2251 | 13.1 | 14.1 | 11.4 | 13.65 | 17.8275 | 13 | 13.4
| Mean |  | 11.0833 | 14.0833 | 14.1833 | 13.1833 | 13.2833 | 12.6042 | 13.05 | 14.35 | 11.3833 | 14.155 | 17.5713 | 12.9167 | 13.7333
| Median |  | 11.05 | 14.1 | 14.3 | 13.2 | 13.35 | 12.6557 | 13.05 | 14.15 | 11.4 | 14.245 | 17.6435 | 13 | 13.6
| Std Dev. |  | 3.1877 | 0.16021 | 0.35449 | 0.2137 | 0.27869 | 0.25726 | 0.10488 | 0.501 | 0.075277 | 0.29494 | 0.27371 | 0.20412 | 0.35024
| Rel.Std.Dev. |  | 28.76% | 1.138% | 2.499% | 1.621% | 2.098% | 2.041% | 0.8037% | 3.491% | 0.6613% | 2.084% | 1.558% | 1.58% | 2.55%
| PDM3 |  | -16.58% | 6.002% | 6.755% | -0.772% | -0.01935% | -5.131% | -1.776% | 8.009% | -14.32% | 6.541% | 32.26% | -2.779% | 3.368%
| Z-Score (Absolute) |  | 0.55 | 0.79 | 0.89 | 0.1 | 0 | 0.67 | 0.23 | 0.89 | 1.89 | 0.87 | 4.26 | 0.28 | 0.31
| Analytical results for Re in OREAS 461 (Certified Value < 0.01 ppm) | [15 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.008 | < 0.002 | 0.004 | < 0.002 | 0.003 | < 0.002 | < 0.02 | < 0.002 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.01 | < 0.002 | 0.005 | < 0.002 | 0.003 | < 0.002 | < 0.02 | 0.002 | 0.003 | < 0.05 | < 0.1 |  |  |  | 17
| 1 | 3 | 0.008 | 0.002 | 0.004 | 0.002 | 0.003 | < 0.002 | < 0.02 | < 0.002 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.008 | 0.003 | 0.003 | 0.005 | 0.003 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1 |  |  |  | < 0.01
| 1 | 5 | 0.007 | 0.006 | 0.004 | 0.003 | 0.004 | < 0.002 | < 0.02 | 0.002 | 0.004 | < 0.05 | < 0.1 |  |  |  | 104
| 1 | 6 | 0.007 | < 0.002 | 0.003 | 0.002 | 0.003 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1
| Mean |  | 01: 0.008; 02: 0.0036667; 03: 0.0038333; 05: 0.003; 06: 0.0031667; 09: 0.0025; 10: 0.0035
| Median |  | 01: 0.008; 02: 0.003; 03: 0.004; 05: 0.0025; 06: 0.003; 09: 0.0025; 10: 0.0035
| Std Dev. |  | 01: 0.0010954; 02: 0.0020817; 03: 0.00075277; 05: 0.0014142; 06: 0.00040825; 09: 0.00057735; 10: 0.00070711
| Rel.Std.Dev. |  | 01: 13.69%; 02: 56.77%; 03: 19.64%; 05: 47.14%; 06: 12.89%; 09: 23.09%; 10: 20.2%
PDM3
| Z-Score (Absolute) |  | 4.59 | 0.59 | 1.21 | 0 | 0.67 | NA | NA | 0.27 | 0.76 | NA | NA
| Analytical results for S in OREAS 461 (Certified Value 475 ppm) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 544 | 400 | 500 | 450 |  |  | 1
| 1 | 2 | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 469 | 300 | 550 | 450 |  |  | 10
| 1 | 3 | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 490 | 548 | 400 | 550 | 500 |  |  | 16
| 1 | 4 | 400 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 546 | 400 | 500 | 500 |  |  | 475.4111
| 1 | 5 | 400 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 551 | 400 | 500 | 500 |  |  | 105
| 1 | 6 | 500 | 500 | 500 | 500 | 500 | 500 | 400 | 480 | 359 | 400 | 550 | 500
| Mean |  | 466.6667 | 500 | 500 | 500 | 500 | 416.6667 | 400 | 481.6667 | 502.8333 | 383.3333 | 525 | 483.3333
| Median |  | 500 | 500 | 500 | 500 | 500 | 400 | 400 | 480 | 545 | 400 | 525 | 500
| Std Dev. |  | 51.6398 | 0 | 0 | 0 | 0 | 40.8248 | 0 | 4.0825 | 77.1373 | 40.8248 | 27.3861 | 25.8199
| Rel.Std.Dev. |  | 11.07% | 0% | 0% | 0% | 0% | 9.798% | 0% | 0.8476% | 15.34% | 10.65% | 5.216% | 5.342%
| PDM3 |  | -1.839% | 5.172% | 5.172% | 5.172% | 5.172% | -12.36% | -15.86% | 1.316% | 5.768% | -19.37% | 10.43% | 1.666%
| Z-Score (Absolute) |  | 0.96 | 0.32 | 0.32 | 0.32 | 0.32 | 2.89 | 3.53 | 0.39 | 1.54 | 3.53 | 1.28 | 0.32
| Analytical results for Sb in OREAS 461 (Certified Value 2.42 ppm) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.35 | 2.67 | 2.58 | 2.49 | 2.43 | 2.1807 | 2.4 | 2.3 | 2.34 | 2.34 | 1.8 |  |  |  | 1
| 1 | 2 | 2.34 | 2.54 | 2.7 | 2.4 | 2.38 | 2.2331 | 2.5 | 2.5 | 2.33 | 2.12 | 1.8 |  |  |  | #N/A
| 1 | 3 | 2.25 | 2.64 | 2.75 | 2.48 | 2.43 | 2.2398 | 2.4 | 2.5 | 2.37 | 2.29 | 2 |  |  |  | 16
| 1 | 4 | 2.51 | 2.68 | 2.74 | 2.65 | 2.35 | 2.386 | 2.5 | 2.1 | 2.34 | 2.1 | 2 |  |  |  | 2.4174
| 1 | 5 | 2.33 | 2.55 | 2.72 | 2.62 | 2.41 | 2.2053 | 2.4 | 2.3 | 2.34 | 2.02 | 2 |  |  |  | 106
| 1 | 6 | 2.31 | 2.52 | 2.65 | 2.62 | 2.48 | 2.4325 | 2.5 | 2.5 | 2.41 | 2.09 | 1.8
| Mean |  | 2.3483 | 2.6 | 2.69 | 2.5433 | 2.4133 | 2.2796 | 2.45 | 2.3667 | 2.355 | 2.16 | 1.9
| Median |  | 2.335 | 2.595 | 2.71 | 2.555 | 2.42 | 2.2364 | 2.45 | 2.4 | 2.34 | 2.11 | 1.9
| Std Dev. |  | 0.086814 | 0.071274 | 0.064498 | 0.10053 | 0.045019 | 0.10367 | 0.054772 | 0.1633 | 0.030166 | 0.1257 | 0.10954
| Rel.Std.Dev. |  | 3.697% | 2.741% | 2.398% | 3.953% | 1.865% | 4.548% | 2.236% | 6.9% | 1.281% | 5.819% | 5.766%
| PDM3 |  | -2.857% | 7.554% | 11.28% | 5.21% | -0.1677% | -5.702% | 1.349% | -2.098% | -2.581% | -10.65% | -21.4%
| Z-Score (Absolute) |  | 0.39 | 1.81 | 2.5 | 1.37 | 0.36 | 0.67 | 0.65 | 0 | 0.09 | 1.6 | 3.61
| Analytical results for Sc in OREAS 461 (Certified Value 38.8 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 1
2
| 1 | 1 | 32.3 | 41.6 | 39.8 | 40.3 | 40.2 | 37.291 | 40 | 36 | 38.2 | 35.2283 | 45 | 38 |  |  | 1
| 1 | 2 | 38.2 | 40.5 | 40.5 | 39.2 | 38 | 37.2491 | 40 | 36 | 38.1 | 35.352 | 44 | 38 |  |  | 18
| 1 | 3 | 39 | 41.5 | 40.8 | 39.9 | 40.3 | 37.8361 | 40 | 37 | 37.5 | 35.1453 | 44 | 39 |  |  | 16
| 1 | 4 | 39.2 | 41 | 39.3 | 39.8 | 38 | 38.2016 | 40 | 37 | 37 | 35.2188 | 44 | 37 |  |  | 38.8492
| 1 | 5 | 35.8 | 41 | 41 | 41.1 | 39.2 | 38.427 | 40 | 37 | 36.4 | 35.3446 | 43 | 38 |  |  | 107
| 1 | 6 | 34 | 41 | 40.1 | 40.9 | 39.1 | 37.9736 | 39 | 37 | 37 | 35.0765 | 44 | 39
| Mean |  | 36.4167 | 41.1 | 40.25 | 40.2 | 39.1333 | 37.8297 | 39.8333 | 36.6667 | 37.3667 | 35.2276 | 44 | 38.1667
| Median |  | 37 | 41 | 40.3 | 40.1 | 39.15 | 37.9049 | 40 | 37 | 37.25 | 35.2235 | 44 | 38
| Std Dev. |  | 2.8555 | 0.4 | 0.64109 | 0.71554 | 1.0073 | 0.47814 | 0.40825 | 0.5164 | 0.70048 | 0.10847 | 0.63246 | 0.75277
| Rel.Std.Dev. |  | 7.841% | 0.9732% | 1.593% | 1.78% | 2.574% | 1.264% | 1.025% | 1.408% | 1.875% | 0.3079% | 1.437% | 1.972%
| PDM3 |  | -6.262% | 5.794% | 3.606% | 3.477% | 0.7313% | -2.624% | 2.533% | -5.618% | -3.816% | -9.322% | 13.26% | -1.757%
| Z-Score (Absolute) |  | 0.96 | 1.05 | 0.69 | 0.66 | 0.21 | 0.35 | 0.51 | 0.85 | 0.55 | 1.47 | 2.29 | 0.21
| Analytical results for Se in OREAS 461 (Indicative Value 6.48 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 4 | < 50 | 5 | 7 | 9 | 0.97792 | 9 | 6 | < 0.5 | < 0.5 | 5 |  |  |  | 1
| 1 | 2 | 4 | < 50 | 6 | 6 | 8 | 0.79493 | 10 | 8 | < 0.5 | < 0.5 | 5 |  |  |  | 15
| 1 | 3 | 4 | < 50 | 6 | 7 | 9 | 0.6892 | 10 | 6 | < 0.5 | < 0.5 | 5 |  |  |  | 16
| 1 | 4 | 4 | < 50 | 5 | 6 | 8 | 0.8 | 10 | 7 | < 0.5 | < 0.5 | 5 |  |  |  | 6.4762
| 1 | 5 | 4 | < 50 | 6 | 6 | 9 | 1.033 | 9 | 6 | < 0.5 | < 0.5 | 5 |  |  |  | 26
| 1 | 6 | 4 | < 50 | 5 | 6 | 8 | 0.86193 | 9 | 6 | < 0.5 | < 0.5 | 5
| Mean |  | 4 |  | 5.5 | 6.3333 | 8.5 | 0.8595 | 9.5 | 6.5 |  |  | 5
| Median |  | 4 |  | 5.5 | 6 | 8.5 | 0.83097 | 9.5 | 6 |  |  | 5
| Std Dev. |  | 0 |  | 0.54772 | 0.5164 | 0.54772 | 0.12716 | 0.54772 | 0.83666 |  |  | 0
| Rel.Std.Dev. |  | 0% |  | 9.959% | 8.154% | 6.444% | 14.79% | 5.766% | 12.87% |  |  | 0%
| PDM3 |  | -38.24% |  | -15.07% | -2.206% | 31.25% | -86.73% | 46.69% | 0.3676% |  |  | -22.79%
| Z-Score (Absolute) |  | 0.91 | NA | 0.2 | 0.2 | 1.23 | 2.41 | 1.71 | 0.28 | NA | NA | 0.44
| Analytical results for Sm in OREAS 461 (Certified Value 208 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 190.5 | 219 | 198.5 | 203 | 202 | 223.7505 | 207.5 | 200.8 | 215 | 206.45 | 213.9425 | 211.6 | 237 |  | 1
| 1 | 2 | 203 | 208 | 205 | 202 | 193.5 | 228.7411 | 204.2 | 212.3 | 216 | 203.43 | 213.105 | 209.8 | 223 |  | 8
| 1 | 3 | 202 | 225 | 208 | 206 | 200 | 221.665 | 209.7 | 206.2 | 214 | 209.18 | 213.4695 | 212.2 | 223 |  | 16
| 1 | 4 | 206 | 221 | 202 | 217 | 195.5 | 233.8646 | 210.7 | 205.5 | 209 | 215.51 | 212.3555 | 211.3 | 219 |  | 208.2366
| 1 | 5 | 196.5 | 210 | 208 | 219 | 197.5 | 233.7077 | 207.2 | 210.5 | 209 | 208.73 | 213.4181 | 212 | 217 |  | 108
| 1 | 6 | 188 | 225 | 206 | 212 | 202 | 228.2028 | 206.8 | 216.2 | 213 | 202.37 | 214.3532 | 215.8 | 235
| Mean |  | 197.6667 | 218 | 204.5833 | 209.8333 | 198.4167 | 228.322 | 207.6833 | 208.5833 | 212.6667 | 207.6117 | 213.4406 | 212.1167 | 225.6667
| Median |  | 199.25 | 220 | 205.5 | 209 | 198.75 | 228.472 | 207.35 | 208.35 | 213.5 | 207.59 | 213.4438 | 211.8 | 223
| Std Dev. |  | 7.2503 | 7.3756 | 3.7204 | 7.2503 | 3.5131 | 5.0026 | 2.2956 | 5.5003 | 3.0111 | 4.7414 | 0.68936 | 1.9944 | 8.3586
| Rel.Std.Dev. |  | 3.668% | 3.383% | 1.819% | 3.455% | 1.771% | 2.191% | 1.105% | 2.637% | 1.416% | 2.284% | 0.323% | 0.9402% | 3.704%
| PDM3 |  | -5.076% | 4.689% | -1.754% | 0.7668% | -4.716% | 9.645% | -0.2657% | 0.1665% | 2.127% | -0.3001% | 2.499% | 1.863% | 8.37%
| Z-Score (Absolute) |  | 2.27 | 1.53 | 0.98 | 0 | 2.13 | 3.46 | 0.4 | 0.23 | 0.53 | 0.42 | 0.67 | 0.43 | 2.96
| Analytical results for Sn in OREAS 461 (Certified Value 21.6 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 20.2 | 23.2 | 24.7 | 22.1 | 21.6 | 20.9812 | 22.8 | 19 | 17.7 | 23.5 | 21 |  |  |  | 1
| 1 | 2 | 20.2 | 22.3 | 25.6 | 21.5 | 21.2 | 19.707 | 22.2 | 19 | 17.3 | 23.4 | 21 |  |  |  | 9
| 1 | 3 | 19.8 | 23.4 | 25.9 | 22.4 | 22 | 19.6587 | 22.7 | 21 | 17.3 | 23.1 | 22 |  |  |  | 16
| 1 | 4 | 20.3 | 23.7 | 25.6 | 22.7 | 21.2 | 20.7433 | 22.1 | 19 | 17.9 | 24.1 | 23 |  |  |  | 21.6377
| 1 | 5 | 20.4 | 22.8 | 25.9 | 23.2 | 21.9 | 21.0801 | 21.5 | 20 | 17.8 | 23.2 | 22 |  |  |  | 109
| 1 | 6 | 20 | 22.9 | 25.3 | 22.8 | 22 | 20.2172 | 22.5 | 20 | 18 | 22.8 | 22
| Mean |  | 20.15 | 23.05 | 25.5 | 22.45 | 21.65 | 20.3979 | 22.3 | 19.6667 | 17.6667 | 23.35 | 21.8333
| Median |  | 20.2 | 23.05 | 25.6 | 22.55 | 21.75 | 20.4802 | 22.35 | 19.5 | 17.75 | 23.3 | 22
| Std Dev. |  | 0.21679 | 0.49295 | 0.45166 | 0.59582 | 0.37815 | 0.6295 | 0.47749 | 0.8165 | 0.30111 | 0.44159 | 0.75277
| Rel.Std.Dev. |  | 1.076% | 2.139% | 1.771% | 2.654% | 1.747% | 3.086% | 2.141% | 4.152% | 1.704% | 1.891% | 3.448%
| PDM3 |  | -6.875% | 6.527% | 17.85% | 3.754% | 0.0569% | -5.73% | 3.061% | -9.109% | -18.35% | 7.914% | 0.9042%
| Z-Score (Absolute) |  | 0.79 | 0.57 | 1.72 | 0.29 | 0.09 | 0.67 | 0.22 | 1.02 | 1.96 | 0.71 | 0
| Analytical results for Sr in OREAS 461 (Certified Value 564 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 601 | 600 | 518 | 576 | 535 | 573.4651 | 559 | 595.1 | 575 | 610.08 | 513.3074 | 369 | 585 | 516 | 1
| 1 | 2 | 619 | 595 | 519 | 581 | 547 | 575.7358 | 564.6 | 599.2 | 566 | 591.47 | 515.5238 | 364 | 585 | 504 | 11
| 1 | 3 | 592 | 602 | 517 | 583 | 520 | 578.9022 | 569.9 | 607 | 574 | 602.58 | 510.9785 | 375 | 585 | 513 | 16
| 1 | 4 | 607 | 603 | 518 | 576 | 501 | 565.6398 | 571.1 | 593.2 | 566 | 589.46 | 515.7509 | 371 | 580 | 508 | 564.1131
| 1 | 5 | 590 | 601 | 515 | 603 | 536 | 559.876 | 557.8 | 604.3 | 553 | 584.08 | 512.1413 | 360 | 590 | 508 | 110
| 1 | 6 | 591 | 605 | 518 | 596 | 505 | 568.3213 | 560.2 | 612 | 572 | 582.61 | 514.5032 | 376 | 580 | 513
| Mean |  | 600 | 601 | 517.5 | 585.8333 | 524 | 570.3234 | 563.7667 | 601.8 | 567.6667 | 593.38 | 513.7009 | 369.1667 | 584.1667 | 510.3333
| Median |  | 596.5 | 601.5 | 518 | 582 | 527.5 | 570.8932 | 562.4 | 601.75 | 569 | 590.465 | 513.9053 | 370 | 585 | 510.5
| Std Dev. |  | 11.4543 | 3.4059 | 1.3784 | 11.1609 | 18.4391 | 7.0406 | 5.712 | 7.2518 | 8.165 | 10.8169 | 1.9062 | 6.2423 | 3.7639 | 4.4121
| Rel.Std.Dev. |  | 1.909% | 0.5667% | 0.2664% | 1.905% | 3.519% | 1.235% | 1.013% | 1.205% | 1.438% | 1.823% | 0.3711% | 1.691% | 0.6443% | 0.8646%
| PDM3 |  | 6.362% | 6.539% | -8.263% | 3.85% | -7.111% | 1.101% | -0.06142% | 6.681% | 0.6299% | 5.188% | -8.937% | -34.56% | 3.555% | -9.534%
| Z-Score (Absolute) |  | 0.66 | 0.69 | 1.1 | 0.36 | 0.96 | 0.03 | 0.11 | 0.7 | 0.03 | 0.52 | 1.18 | 4.28 | 0.32 | 1.26
| Analytical results for Ta in OREAS 461 (Certified Value 20.9 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 22.9 | 23.6 | 21.9 | 21.8 | 18.6525 | 21.6 | 12.3 | 11.1 | 20.87 | 19.1 |  |  |  |  | 1
| 1 | 2 | 23.5 | 23 | 20.9 | 21.6 | 17.902 | 21.5 | 13.1 | 11.1 | 15.5 | 19.3 |  |  |  |  | 1
| 1 | 3 | 22.7 | 23.7 | 21.8 | 22.5 | 18.6311 | 22.6 | 10.8 | 10.8 | 17.62 | 21.2 |  |  |  |  | 16
| 1 | 4 | 22.7 | 22 | 22.1 | 21.5 | 18.8859 | 21 | 13 | 11.3 | 17.48 | 19.7 |  |  |  |  | 20.9304
| 1 | 5 | 22.2 | 23.5 | 22.5 | 22 | 20.3728 | 20.9 | 12.5 | 11.5 | 17.92 | 20.7 |  |  |  |  | 111
| 1 | 6 | 22.7 | 23 | 22.4 | 22.3 | 19.4891 | 20.7 | 12.5 | 11.5 | 17.75 | 20.1
| Mean |  | 22.7833 | 23.1333 | 21.9333 | 21.95 | 18.9889 | 21.3833 | 12.3667 | 11.2167 | 17.8567 | 20.0167
| Median |  | 22.7 | 23.25 | 22 | 21.9 | 18.7692 | 21.25 | 12.5 | 11.2 | 17.685 | 19.9
| Std Dev. |  | 0.42151 | 0.6314 | 0.57504 | 0.3937 | 0.84834 | 0.69113 | 0.82865 | 0.27142 | 1.7232 | 0.81588
| Rel.Std.Dev. |  | 1.85% | 2.729% | 2.622% | 1.794% | 4.468% | 3.232% | 6.701% | 2.42% | 9.65% | 4.076%
| PDM3 |  | 8.853% | 10.53% | 4.792% | 4.872% | -9.276% | 2.164% | -40.92% | -46.41% | -14.69% | -4.365%
| Z-Score (Absolute) |  | 0.74 | 0.86 | 0.44 | 0.44 | 0.61 | 0.24 | 2.85 | 3.37 | 1.22 | 0.24
| Analytical results for Tb in OREAS 461 (Certified Value 9.09 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 8.67 | 8.48 | 10.2 | 8.62 | 10 | 9.0451 | 9.1 | 8.75 | 11.2 | 9.59 | 13.8753 | 8.9 | 8.6 |  | 1
| 1 | 2 | 9.46 | 9.05 | 9.62 | 8.63 | 9.53 | 9.1344 | 9.1 | 9.08 | 11.1 | 9.3 | 13.9624 | 8.75 | 7.96 |  | 2
| 1 | 3 | 9.52 | 8.73 | 9.69 | 8.88 | 10.05 | 9.2887 | 9.3 | 8.99 | 11.1 | 9.4 | 14.2096 | 8.95 | 8.02 |  | 16
| 1 | 4 | 9.75 | 9.01 | 8.35 | 8.67 | 9.52 | 9.1004 | 9.2 | 8.7 | 11.1 | 9.7 | 14.1303 | 8.8 | 8.04 |  | 9.0942
| 1 | 5 | 9.41 | 8.62 | 9.77 | 8.7 | 10.1 | 9.2158 | 9.1 | 9.07 | 11.1 | 9.39 | 14.0595 | 8.75 | 8 |  | 112
| 1 | 6 | 9.09 | 8.78 | 9.32 | 9.05 | 9.93 | 9.0506 | 9.3 | 9.02 | 11.2 | 9.27 | 14.1682 | 9.05 | 8.48
| Mean |  | 9.3167 | 8.7783 | 9.4917 | 8.7583 | 9.855 | 9.1392 | 9.1833 | 8.935 | 11.1333 | 9.4417 | 14.0676 | 8.8667 | 8.1833
| Median |  | 9.435 | 8.755 | 9.655 | 8.685 | 9.965 | 9.1174 | 9.15 | 9.005 | 11.1 | 9.395 | 14.0949 | 8.85 | 8.03
| Std Dev. |  | 0.38156 | 0.22085 | 0.62736 | 0.17128 | 0.26174 | 0.096405 | 0.098319 | 0.1667 | 0.05164 | 0.16893 | 0.12831 | 0.12111 | 0.28012
| Rel.Std.Dev. |  | 4.095% | 2.516% | 6.61% | 1.956% | 2.656% | 1.055% | 1.071% | 1.866% | 0.4638% | 1.789% | 0.9121% | 1.366% | 3.423%
| PDM3 |  | 2.447% | -3.473% | 4.371% | -3.693% | 8.366% | 0.495% | 0.9805% | -1.75% | 22.42% | 3.821% | 54.69% | -2.502% | -10.02%
| Z-Score (Absolute) |  | 0.22 | 0.67 | 0.89 | 0.8 | 1.12 | 0.07 | 0 | 0.41 | 3.25 | 0.43 | 8.13 | 0.53 | 1.8
| Analytical results for Te in OREAS 461 (Certified Value 0.34 ppm) | [15 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.31 | 0.33 | 0.39 | 0.32 | 0.27 | 0.29853 | < 1 | < 0.5 | 0.37 | 0.3 | 0.2 |  |  |  | 1
| 1 | 2 | 0.26 | 0.37 | 0.44 | 0.32 | 0.31 | 0.40095 | < 1 | < 0.5 | 0.39 | 0.3 | 0.2 |  |  |  | 16
| 1 | 3 | 0.27 | 0.33 | 0.39 | 0.31 | 0.32 | 0.29853 | < 1 | < 0.5 | 0.42 | 0.3 | 0.2 |  |  |  | 16
| 1 | 4 | 0.27 | 0.29 | 0.49 | 0.31 | 0.3 | 0.3295 | < 1 | < 0.5 | 0.42 | 0.3 | 0.2 |  |  |  | 0.34127
| 1 | 5 | 0.31 | 0.36 | 0.44 | 0.35 | 0.31 | 0.40241 | < 1 | < 0.5 | 0.4 | 0.3 | 0.2 |  |  |  | 113
| 1 | 6 | 0.3 | 0.31 | 0.4 | 0.33 | 0.29 | 0.441 | < 1 | < 0.5 | 0.41 | 0.3 | 0.2
| Mean |  | 0.28667 | 0.33167 | 0.425 | 0.32333 | 0.3 | 0.36182 |  |  | 0.40167 | 0.3 | 0.2
| Median |  | 0.285 | 0.33 | 0.42 | 0.32 | 0.305 | 0.36522 |  |  | 0.405 | 0.3 | 0.2
| Std Dev. |  | 0.022509 | 0.029944 | 0.03937 | 0.015055 | 0.017889 | 0.060833 |  |  | 0.019408 | 0 | 3.0405e-17
| Rel.Std.Dev. |  | 7.852% | 9.028% | 9.264% | 4.656% | 5.963% | 16.81% |  |  | 4.832% | 0% | 1.52e-14%
| PDM3 |  | -16% | -2.814% | 24.54% | -5.256% | -12.09% | 6.022% |  |  | 17.7% | -12.09% | -41.4%
| Z-Score (Absolute) |  | 0.67 | 0.15 | 1.87 | 0 | 0.43 | 0.71 | NA | NA | 1.44 | 0.43 | 2.27
| Analytical results for Th in OREAS 461 (Certified Value 207 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 172.5 | 195 | 216 | 188.5 | 208 | 197.5202 | 199 | 215 | 196.86 | 214.71 | 198 |  |  |  | 1
| 1 | 2 | 210 | 212 | 223 | 189 | 199 | 197.4394 | 204.8 | 219 | 196.22 | 216.5911 | 201 |  |  |  | #N/A
| 1 | 3 | 216 | 214 | 225 | 193 | 212 | 200.1387 | 202.3 | 219 | 203.54 | 215.0865 | 203 |  |  |  | 16
| 1 | 4 | 219 | 211 | 218 | 198.5 | 202 | 204.8934 | 201.2 | 210 | 204.43 | 215.1762 | 203 |  |  |  | 206.55
| 1 | 5 | 187 | 210 | 225 | 197.5 | 204 | 204.727 | 205.9 | 215 | 199.4 | 214.8293 | 203 |  |  |  | 114
| 1 | 6 | 181.5 | 209 | 221 | 204 | 207 | 199.2399 | 207 | 216 | 198.33 | 215.4659 | 213
| Mean |  | 197.6667 | 208.5 | 221.3333 | 195.0833 | 205.3333 | 200.6598 | 203.3667 | 215.6667 | 199.7967 | 215.3098 | 203.5
| Median |  | 198.5 | 210.5 | 222 | 195.25 | 205.5 | 199.6893 | 203.55 | 215.5 | 198.865 | 215.1314 | 203
| Std Dev. |  | 19.7577 | 6.8337 | 3.7238 | 6.0284 | 4.6332 | 3.3758 | 3.052 | 3.3267 | 3.4411 | 0.68183 | 5.0498
| Rel.Std.Dev. |  | 9.995% | 3.278% | 1.682% | 3.09% | 2.256% | 1.682% | 1.501% | 1.543% | 1.722% | 0.3167% | 2.481%
| PDM3 |  | -4.301% | 0.9441% | 7.157% | -5.552% | -0.589% | -2.852% | -1.541% | 4.414% | -3.27% | 4.241% | -1.477%
| Z-Score (Absolute) |  | 0.13 | 1.48 | 3.39 | 1.56 | 0.37 | 0.51 | 0 | 2.32 | 0.67 | 2.26 | 0.33
| Analytical results for Ti in OREAS 461 (Certified Value 1.06 wt.%) | [15 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 1.12 | 1.155 | 1.1 | 1.015 | 1.005 | 1.032 | 1.03 | 0.6646 | 0.848 | 1.2764 | 0.66 | 1.33 |  |  | 1
| 1 | 2 | 1.17 | 1.125 | 1.085 | 0.973 | 1.07 | 0.95489 | 0.99 | 0.6268 | 0.881 | 1.2904 | 0.66 | 1.32 |  |  | 3
| 1 | 3 | 1.135 | 1.17 | 1.065 | 1.02 | 1.035 | 0.98844 | 1.01 | 0.6967 | 0.877 | 1.2475 | 0.65 | 1.45 |  |  | 16
| 1 | 4 | 1.105 | 1.24 | 1.08 | 1 | 0.991 | 0.93209 | 0.94 | 0.6788 | 0.856 | 1.2754 | 0.66 | 1.42 |  |  | 1.0566
| 1 | 5 | 1.095 | 1.19 | 1.085 | 1.04 | 1.065 | 0.95876 | 0.93 | 0.689 | 0.854 | 1.2401 | 0.67 | 1.4 |  |  | 115
| 1 | 6 | 1.135 | 1.175 | 1.105 | 1.06 | 1.015 | 1.022 | 0.93 | 0.6851 | 0.851 | 1.216 | 0.67 | 1.38
| Mean |  | 1.1267 | 1.1758 | 1.0867 | 1.018 | 1.0302 | 0.98137 | 0.97167 | 0.6735 | 0.86117 | 1.2576 | 0.66167 | 1.3833
| Median |  | 1.1275 | 1.1725 | 1.085 | 1.0175 | 1.025 | 0.9736 | 0.965 | 0.68195 | 0.855 | 1.2614 | 0.66 | 1.39
| Std Dev. |  | 0.026583 | 0.038395 | 0.014376 | 0.030364 | 0.032314 | 0.039773 | 0.044008 | 0.025303 | 0.014134 | 0.027874 | 0.0075277 | 0.050859
| Rel.Std.Dev. |  | 2.359% | 3.265% | 1.323% | 2.983% | 3.137% | 4.053% | 4.529% | 3.757% | 1.641% | 2.216% | 1.138% | 3.677%
| PDM3 |  | 6.634% | 11.29% | 2.848% | -3.651% | -2.499% | -7.118% | -8.036% | -36.26% | -18.49% | 19.03% | -37.38% | 30.93%
| Z-Score (Absolute) |  | 0.54 | 0.8 | 0.33 | 0.03 | 0.03 | 0.23 | 0.28 | 1.81 | 0.86 | 1.24 | 1.92 | 1.9
| Analytical results for Tl in OREAS 461 (Certified Value 0.1 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.11 | 0.16 | 0.11 | 0.12 | 0.11 | 0.097011 | 0.1 | < 0.2 | 0.1 | 0.1 | 0.22592 | 0.1 | < 0.1 |  | 1
| 1 | 2 | 0.12 | 0.21 | 0.1 | 0.11 | 0.11 | 0.095978 | 0.12 | < 0.2 | 0.11 | 0.1 | 0.20547 | 0.1 | 0.1 |  | 17
| 1 | 3 | 0.11 | 0.23 | 0.11 | 0.1 | 0.11 | 0.10073 | 0.1 | < 0.2 | 0.1 | 0.1 | 0.22987 | 0.1 | < 0.1 |  | 16
| 1 | 4 | 0.12 | 0.22 | 0.1 | 0.1 | 0.09 | 0.11122 | 0.11 | < 0.2 | 0.1 | 0.09 | 0.19515 | 0.1 | < 0.1 |  | 0.10384
| 1 | 5 | 0.11 | 0.21 | 0.11 | 0.11 | 0.11 | 0.11527 | 0.1 | < 0.2 | 0.1 | 0.09 | 0.2047 | 0.1 | < 0.1 |  | 116
| 1 | 6 | 0.11 | 0.19 | 0.11 | 0.11 | 0.1 | 0.11994 | 0.08 | < 0.2 | 0.1 | 0.09 | 0.18985 | 0.1 | 0.1
| Mean |  | 0.11333 | 0.20333 | 0.10667 | 0.10833 | 0.105 | 0.10669 | 0.10167 |  | 0.10167 | 0.095 | 0.20849 | 0.1 | 0.1
| Median |  | 0.11 | 0.21 | 0.11 | 0.11 | 0.11 | 0.10598 | 0.1 |  | 0.1 | 0.095 | 0.20509 | 0.1 | 0.1
| Std Dev. |  | 0.005164 | 0.025033 | 0.005164 | 0.0075277 | 0.0083666 | 0.010137 | 0.013292 |  | 0.0040825 | 0.0054772 | 0.016183 | 1.5202e-17 | 0
| Rel.Std.Dev. |  | 4.556% | 12.31% | 4.841% | 6.949% | 7.968% | 9.501% | 13.07% |  | 4.016% | 5.766% | 7.762% | 1.52e-14% | 0%
| PDM3 |  | 9.147% | 95.82% | 2.726% | 4.331% | 1.121% | 2.751% | -2.089% |  | -2.089% | -8.51% | 100.8% | -3.694% | -3.694%
| Z-Score (Absolute) |  | 1.01 | 13.15 | 0.11 | 0.34 | 0.11 | 0.12 | 0.56 | NA | 0.56 | 1.46 | 13.85 | 0.79 | 0.79
| Analytical results for Tm in OREAS 461 (Certified Value 0.7 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.75 | 0.72 | 0.74 | 0.68 | 0.72 | 0.7271 | 0.6543 | 0.8 | 0.66 | 0.61 | 0.85699 | 0.65 | 0.7 |  | 1
| 1 | 2 | 0.82 | 0.71 | 0.74 | 0.68 | 0.69 | 0.75963 | 0.69235 | 0.7 | 0.67 | 0.63 | 0.84328 | 0.65 | 0.64 |  | 4
| 1 | 3 | 0.83 | 0.73 | 0.74 | 0.7 | 0.72 | 0.71593 | 0.69481 | 0.7 | 0.67 | 0.63 | 0.85542 | 0.65 | 0.7 |  | 16
| 1 | 4 | 0.81 | 0.73 | 0.67 | 0.69 | 0.7 | 0.71171 | 0.69522 | 0.7 | 0.66 | 0.65 | 0.85063 | 0.65 | 0.7 |  | 0.70061
| 1 | 5 | 0.77 | 0.73 | 0.75 | 0.69 | 0.72 | 0.73368 | 0.64421 | 0.7 | 0.67 | 0.65 | 0.87224 | 0.65 | 0.68 |  | 117
| 1 | 6 | 0.76 | 0.72 | 0.71 | 0.7 | 0.74 | 0.75799 | 0.69714 | 0.7 | 0.67 | 0.61 | 0.89374 | 0.65 | 0.7
| Mean |  | 0.79 | 0.72333 | 0.725 | 0.69 | 0.715 | 0.73434 | 0.67967 | 0.71667 | 0.66667 | 0.63 | 0.86205 | 0.65 | 0.68667
| Median |  | 0.79 | 0.725 | 0.74 | 0.69 | 0.72 | 0.73039 | 0.69358 | 0.7 | 0.67 | 0.63 | 0.8562 | 0.65 | 0.7
| Std Dev. |  | 0.034059 | 0.008165 | 0.030166 | 0.0089443 | 0.017607 | 0.020507 | 0.023825 | 0.040825 | 0.005164 | 0.017889 | 0.018224 | 0 | 0.024221
| Rel.Std.Dev. |  | 4.311% | 1.129% | 4.161% | 1.296% | 2.462% | 2.793% | 3.505% | 5.696% | 0.7746% | 2.839% | 2.114% | 0% | 3.527%
| PDM3 |  | 12.76% | 3.243% | 3.481% | -1.515% | 2.054% | 4.814% | -2.988% | 2.292% | -4.845% | -10.08% | 23.04% | -7.224% | -1.99%
| Z-Score (Absolute) |  | 1.78 | 0.2 | 0.24 | 0.59 | 0 | 0.46 | 0.84 | 0.04 | 1.15 | 2.02 | 3.5 | 1.55 | 0.67
| Analytical results for U in OREAS 461 (Certified Value 4.4 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.4 | 4.2 | 4.5 | 4.2 | 4.5 | 4.2287 | 4.41 | 4.1 | 4.37 | 4.32 | 5.0784 | 4.1 | 3.7 |  | 1
| 1 | 2 | 4.4 | 4.4 | 4.6 | 4.2 | 4.4 | 4.2613 | 4.37 | 4.2 | 4.44 | 4.17 | 4.7527 | 4 | 3.8 |  | 5
| 1 | 3 | 4.4 | 4.4 | 4.8 | 4.3 | 4.6 | 4.4315 | 4.38 | 4.1 | 4.44 | 4.34 | 5.0481 | 4.1 | 3.8 |  | 16
| 1 | 4 | 4.5 | 4.5 | 4.6 | 4.5 | 4.4 | 4.4749 | 4.23 | 4 | 4.38 | 4.43 | 4.8021 | 4.1 | 3.8 |  | 4.4013
| 1 | 5 | 4.3 | 4.4 | 4.7 | 4.4 | 4.5 | 4.6947 | 4.41 | 4.2 | 4.55 | 4.62 | 4.973 | 4 | 3.8 |  | 118
| 1 | 6 | 4.5 | 4.3 | 4.7 | 4.5 | 4.5 | 4.5023 | 4.39 | 4.3 | 4.45 | 4.12 | 5.1162 | 4.2 | 4.1
| Mean |  | 4.4167 | 4.3667 | 4.65 | 4.35 | 4.4833 | 4.4322 | 4.365 | 4.15 | 4.4383 | 4.3333 | 4.9618 | 4.0833 | 3.8333
| Median |  | 4.4 | 4.4 | 4.65 | 4.35 | 4.5 | 4.4532 | 4.385 | 4.15 | 4.44 | 4.33 | 5.0106 | 4.1 | 3.8
| Std Dev. |  | 0.075277 | 0.10328 | 0.10488 | 0.13784 | 0.075277 | 0.17105 | 0.068044 | 0.10488 | 0.064317 | 0.18107 | 0.15116 | 0.075277 | 0.13663
| Rel.Std.Dev. |  | 1.704% | 2.365% | 2.256% | 3.169% | 1.679% | 3.859% | 1.559% | 2.527% | 1.449% | 4.179% | 3.046% | 1.844% | 3.564%
| PDM3 |  | 0.3501% | -0.7859% | 5.652% | -1.165% | 1.865% | 0.7038% | -0.8238% | -5.709% | 0.8424% | -1.543% | 12.74% | -7.223% | -12.9%
| Z-Score (Absolute) |  | 0.28 | 0.29 | 2.97 | 0.48 | 1.05 | 0.46 | 0 | 2.78 | 0.53 | 0.67 | 6.55 | 3.55 | 7.03
| Analytical results for V in OREAS 461 (Certified Value 352 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 341 | 372 | 297 | 326 | 310 | 353.9416 | 342.7 | 363 | 352 | 392 | 375.88 | 165 | 370 |  | 1
| 1 | 2 | 340 | 370 | 299 | 330 | 322 | 355.6957 | 354.2 | 361 | 356 | 402 | 374.834 | 164 | 370 |  | 6
| 1 | 3 | 330 | 376 | 294 | 329 | 309 | 353.0633 | 354 | 367 | 351 | 397 | 376.89 | 162 | 380 |  | 16
| 1 | 4 | 327 | 379 | 298 | 328 | 299 | 372.3048 | 366.4 | 364 | 351 | 410 | 374.988 | 164 | 360 |  | 352.0608
| 1 | 5 | 326 | 377 | 300 | 337 | 318 | 367.734 | 366.7 | 365 | 352 | 394 | 376.861 | 161 | 365 |  | 119
| 1 | 6 | 332 | 372 | 300 | 336 | 297 | 366.6284 | 366.9 | 370 | 341 | 394 | 375.654 | 165 | 375
| Mean |  | 332.6667 | 374.3333 | 298 | 331 | 309.1667 | 361.5613 | 358.4833 | 365 | 350.5 | 398.1667 | 375.8512 | 163.5 | 370
| Median |  | 331 | 374 | 298.5 | 329.5 | 309.5 | 361.1621 | 360.3 | 364.5 | 351.5 | 395.5 | 375.767 | 164 | 370
| Std Dev. |  | 6.4395 | 3.5024 | 2.2804 | 4.4721 | 9.9482 | 8.2931 | 9.8852 | 3.1623 | 5.01 | 6.7651 | 0.88525 | 1.6432 | 7.0711
| Rel.Std.Dev. |  | 1.936% | 0.9356% | 0.7652% | 1.351% | 3.218% | 2.294% | 2.758% | 0.8664% | 1.429% | 1.699% | 0.2355% | 1.005% | 1.911%
| PDM3 |  | -5.509% | 6.326% | -15.36% | -5.982% | -12.18% | 2.699% | 1.824% | 3.675% | -0.4433% | 13.1% | 6.757% | -53.56% | 5.095%
| Z-Score (Absolute) |  | 1 | 0.62 | 2.35 | 1.07 | 1.91 | 0.12 | 0 | 0.25 | 0.31 | 1.54 | 0.67 | 7.57 | 0.45
| Analytical results for W in OREAS 461 (Certified Value 2.5 ppm) | [15 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.6 | 2.7 | 3.3 | 2.7 | 2.8 | 2.4648 | 2 | 2 | 2.2 | 2.4 | 2.5 |  |  |  | 1
| 1 | 2 | 2.9 | 4 | 3.2 | 2.7 | 2.7 | 2.2734 | 2 | 2 | 2.1 | 2.2 | 2.5 |  |  |  | 7
| 1 | 3 | 2.7 | 2.7 | 3.3 | 2.7 | 2.8 | 2.3928 | 2 | 2 | 2.2 | 2.4 | 2.5 |  |  |  | 16
| 1 | 4 | 2.8 | 2.8 | 3 | 2.7 | 2.7 | 2.4267 | 2 | 2 | 2.4 | 2.2 | 2.5 |  |  |  | 2.4971
| 1 | 5 | 2.6 | 2.7 | 3.2 | 2.8 | 2.8 | 2.3948 | 2 | 2 | 2.2 | 2.2 | 2.5 |  |  |  | 120
| 1 | 6 | 2.7 | 2.7 | 3.1 | 2.8 | 2.8 | 2.3342 | 2 | 2 | 2.3 | 2.2 | 2.5
| Mean |  | 2.7167 | 2.9333 | 3.1833 | 2.7333 | 2.7667 | 2.3811 | 2 | 2 | 2.2333 | 2.2667 | 2.5
| Median |  | 2.7 | 2.7 | 3.2 | 2.7 | 2.8 | 2.3938 | 2 | 2 | 2.2 | 2.2 | 2.5
| Std Dev. |  | 0.1169 | 0.52409 | 0.1169 | 0.05164 | 0.05164 | 0.068145 | 0 | 0 | 0.10328 | 0.10328 | 0
| Rel.Std.Dev. |  | 4.303% | 17.87% | 3.672% | 1.889% | 1.866% | 2.862% | 0% | 0% | 4.624% | 4.556% | 0%
| PDM3 |  | 8.794% | 17.47% | 27.48% | 9.462% | 10.8% | -4.644% | -19.91% | -19.91% | -10.56% | -9.227% | 0.1173%
| Z-Score (Absolute) |  | 0.63 | 0.64 | 1.97 | 0.67 | 0.77 | 0.34 | 1.44 | 1.44 | 0.87 | 0.67 | 0
| Analytical results for Y in OREAS 461 (Certified Value 83 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 76.3 | 83 | 89.3 | 86.4 | 86.7 | 78.6506 | 83.9 | 80.8 | 81.5 | 87.06 | 84.3177 | 77.1 | 79.7 |  | 1
| 1 | 2 | 84.8 | 77.1 | 87.5 | 85 | 82.1 | 79.3019 | 84 | 82.1 | 79.5 | 87.87 | 84.2674 | 76.7 | 84.4 |  | 8
| 1 | 3 | 85.9 | 80.8 | 90.1 | 86.9 | 86 | 83.9277 | 85.4 | 81.8 | 81.9 | 87.79 | 83.6995 | 77.4 | 81.4 |  | 16
| 1 | 4 | 89.6 | 80.2 | 85.2 | 88.2 | 82.7 | 83.3788 | 88.6 | 80 | 79.9 | 85.96 | 84.2908 | 77.4 | 80.7 |  | 83.097
| 1 | 5 | 80.6 | 77.6 | 90.5 | 89.1 | 85.9 | 80.175 | 84.8 | 82 | 78.8 | 85.72 | 85.1148 | 76.9 | 82 |  | 121
| 1 | 6 | 78.9 | 78.1 | 84.4 | 88.3 | 86.3 | 78.9576 | 86 | 83.5 | 79.5 | 84.62 | 85.9227 | 77.7 | 101
| Mean |  | 82.6833 | 79.4667 | 87.8333 | 87.3167 | 84.95 | 80.7319 | 85.45 | 81.7 | 80.1833 | 86.5033 | 84.6022 | 77.2 | 84.8667
| Median |  | 82.7 | 79.15 | 88.4 | 87.55 | 85.95 | 79.7384 | 85.1 | 81.9 | 79.7 | 86.51 | 84.3042 | 77.25 | 81.7
| Std Dev. |  | 4.941 | 2.2695 | 2.5781 | 1.5039 | 2.0037 | 2.3261 | 1.741 | 1.2 | 1.2336 | 1.2876 | 0.78882 | 0.36878 | 8.0597
| Rel.Std.Dev. |  | 5.976% | 2.856% | 2.935% | 1.722% | 2.359% | 2.881% | 2.037% | 1.469% | 1.538% | 1.488% | 0.9324% | 0.4777% | 9.497%
| PDM3 |  | -0.4978% | -4.369% | 5.7% | 5.078% | 2.23% | -2.846% | 2.832% | -1.681% | -3.506% | 4.099% | 1.811% | -7.097% | 2.13%
| Z-Score (Absolute) |  | 0 | 0.87 | 1.39 | 1.25 | 0.61 | 0.53 | 0.75 | 0.27 | 0.67 | 1.03 | 0.52 | 1.48 | 0.28
| Analytical results for Yb in OREAS 461 (Certified Value 3.06 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.09 | 2.95 | 3.06 | 2.88 | 3.02 | 2.9613 | 2.9006 | 2.9 | 3.81 | 3.02 | 3.4615 | 3 | 3.2 |  | 1
| 1 | 2 | 3.36 | 2.97 | 3.12 | 2.89 | 2.94 | 2.9617 | 2.8247 | 2.9 | 3.84 | 3.17 | 3.2961 | 3 | 3.35 |  | 9
| 1 | 3 | 3.34 | 3.01 | 3.05 | 2.95 | 3.06 | 3.3061 | 2.9131 | 3 | 3.81 | 3.04 | 3.5337 | 3.05 | 3.3 |  | 16
| 1 | 4 | 3.36 | 3.01 | 3.04 | 2.93 | 2.95 | 3.1299 | 3.0184 | 3 | 3.75 | 3.18 | 3.4511 | 2.95 | 3.25 |  | 3.0631
| 1 | 5 | 3.18 | 3.01 | 3.16 | 2.92 | 3.02 | 3.33 | 2.9369 | 2.9 | 3.79 | 3.15 | 3.6928 | 3.05 | 3.3 |  | 122
| 1 | 6 | 3.07 | 2.98 | 3.06 | 3 | 3.08 | 3.3169 | 2.8287 | 3.1 | 3.92 | 3.95 | 2.8579 | 3 | 3.45
| Mean |  | 3.2333 | 2.9883 | 3.0817 | 2.9283 | 3.0117 | 3.1677 | 2.9037 | 2.9667 | 3.82 | 3.2517 | 3.3822 | 3.0083 | 3.3083
| Median |  | 3.26 | 2.995 | 3.06 | 2.925 | 3.02 | 3.218 | 2.9069 | 2.95 | 3.81 | 3.16 | 3.4563 | 3 | 3.3
| Std Dev. |  | 0.13677 | 0.025626 | 0.047504 | 0.043551 | 0.056716 | 0.1756 | 0.072409 | 0.08165 | 0.057271 | 0.34879 | 0.28734 | 0.037639 | 0.08612
| Rel.Std.Dev. |  | 4.23% | 0.8575% | 1.542% | 1.487% | 1.883% | 5.544% | 2.494% | 2.752% | 1.499% | 10.73% | 8.496% | 1.251% | 2.603%
| PDM3 |  | 5.557% | -2.442% | 0.6052% | -4.401% | -1.68% | 3.413% | -5.203% | -3.149% | 24.71% | 6.155% | 10.42% | -1.789% | 8.005%
| Z-Score (Absolute) |  | 1.11 | 0.52 | 0 | 0.91 | 0.36 | 0.67 | 1.08 | 0.66 | 5 | 0.31 | 2.79 | 0.38 | 1.61
| Analytical results for Zn in OREAS 461 (Certified Value 139 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 138 | 142 | 119 | 120 | 119 | 147.3384 | 145.4541 | 142 | 156 | 148 | 222.899 | 85 | 145 | 142 | 1
| 1 | 2 | 137 | 141 | 119 | 120 | 124 | 151.7959 | 145.1305 | 153 | 157 | 152 | 223.209 | 87 | 145 | 132 | #N/A
| 1 | 3 | 133 | 143 | 117 | 122 | 120 | 148.195 | 145.7471 | 171 | 157 | 145 | 223.033 | 87 | 145 | 144 | 16
| 1 | 4 | 133 | 147 | 118 | 122 | 115 | 149.7978 | 144.214 | 145 | 157 | 147 | 221.409 | 89 | 145 | 132 | 139.3186
| 1 | 5 | 131 | 143 | 119 | 126 | 124 | 152.1224 | 143.1699 | 150 | 158 | 145 | 221.007 | 87 | 140 | 138 | 123
| 1 | 6 | 133 | 142 | 119 | 125 | 115 | 148.1167 | 144.6549 | 163 | 158 | 143 | 222.808 | 88 | 145 | 144
| Mean |  | 134.1667 | 143 | 118.5 | 122.5 | 119.5 | 149.561 | 144.7284 | 154 | 157.1667 | 146.6667 | 222.3942 | 87.1667 | 144.1667 | 138.6667
| Median |  | 133 | 142.5 | 119 | 122 | 119.5 | 148.9964 | 144.8927 | 151.5 | 157 | 146 | 222.8535 | 87 | 145 | 140
| Std Dev. |  | 2.7142 | 2.0976 | 0.83666 | 2.51 | 4.0373 | 2.025 | 0.94074 | 11.0635 | 0.75277 | 3.1411 | 0.93733 | 1.3292 | 2.0412 | 5.6095
| Rel.Std.Dev. |  | 2.023% | 1.467% | 0.706% | 2.049% | 3.379% | 1.354% | 0.65% | 7.184% | 0.479% | 2.142% | 0.4215% | 1.525% | 1.416% | 4.045%
| PDM3 |  | -3.698% | 2.642% | -14.94% | -12.07% | -14.23% | 7.352% | 3.883% | 10.54% | 12.81% | 5.274% | 59.63% | -37.43% | 3.48% | -0.4679%
| Z-Score (Absolute) |  | 0.61 | 0.07 | 1.68 | 1.41 | 1.61 | 0.43 | 0.11 | 0.74 | 0.95 | 0.24 | 5.39 | 3.81 | 0.07 | 0.31
| Analytical results for Zr in OREAS 461 (Certified Value 270 ppm) | [15 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 291 | 287 | 272 | 264 | 284 | 244.2632 | 219 | 325 | 264.2 | 194 | 245 | 284 |  |  | 1
| 1 | 2 | 309 | 289 | 265 | 256 | 279 | 242.4219 | 210 | 321 | 271.6 | 194 | 240 | 289 |  |  | 10
| 1 | 3 | 300 | 276 | 301 | 264 | 276 | 247.4138 | 216 | 339 | 263.6 | 209 | 240 | 282 |  |  | 16
| 1 | 4 | 292 | 292 | 261 | 263 | 267 | 265.4465 | 214 | 324 | 258.2 | 200 | 230 | 280 |  |  | 269.8554
| 1 | 5 | 287 | 274 | 308 | 268 | 277 | 265.4442 | 220 | 327 | 253.9 | 180 | 235 | 276 |  |  | 124
| 1 | 6 | 290 | 278 | 265 | 268 | 275 | 254.7995 | 217 | 341 | 258.9 | 194 | 240 | 280
| Mean |  | 294.8333 | 282.6667 | 278.6667 | 263.8333 | 276.3333 | 253.2982 | 216 | 329.5 | 261.7333 | 195.1667 | 238.3333 | 281.8333
| Median |  | 291.5 | 282.5 | 268.5 | 264 | 276.5 | 251.1067 | 216.5 | 326 | 261.25 | 194 | 240 | 281
| Std Dev. |  | 8.1833 | 7.5807 | 20.4418 | 4.4008 | 5.5737 | 10.3142 | 3.6332 | 8.3845 | 6.1448 | 9.4745 | 5.164 | 4.4008
| Rel.Std.Dev. |  | 2.776% | 2.682% | 7.336% | 1.668% | 2.017% | 4.072% | 1.682% | 2.545% | 2.348% | 4.855% | 2.167% | 1.561%
| PDM3 |  | 9.256% | 4.747% | 3.265% | -2.232% | 2.401% | -6.136% | -19.96% | 22.1% | -3.01% | -27.68% | -11.68% | 4.439%
| Z-Score (Absolute) |  | 1.01 | 0.58 | 0.39 | 0.29 | 0.29 | 0.77 | 2.48 | 2.73 | 0.38 | 3.44 | 1.46 | 0.54
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.