## Abbreviations
Table 1. Abbreviations used for OREAS 464
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 464
Lab Slot | Laboratory Name and Location
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
Nagrom, Perth, WA, Australia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
SGS Vostok Limited, Chita, Russian Federation
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 464
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate Fusion XRF
CeO2, Cerium(IV) oxide (wt.%) | 1.895 | 0.027208 | 1.868 | 1.922 | 1.8722 | 1.9178
Fe2O3, Iron(III) oxide (wt.%) | 54.381 | 1.031 | 52.9889 | 55.7731 | 53.9376 | 54.8244
La2O3, Lanthanum(III) oxide (wt.%) | 1.4064 | 0.019739 | 1.3895 | 1.4232 | 1.3847 | 1.428
Nd2O3, Neodymium(III) oxide (wt.%) | 1.124 | 0.036137 | 1.0849 | 1.1632 | 1.1018 | 1.1462
Pr6O11, Praseodymium(III,IV) oxide (ppm) | 3069.6676 | 98.5456 | 2987.2442 | 3152.091 | 2907.4575 | 3231.8776
Sm2O3, Samarium(III) oxide (ppm) | 1634.7163 | 55.2316 | 1583.9556 | 1685.4769 | IND | IND
ThO2, Thorium dioxide (ppm) | 575.7804 | 84.4666 | 491.5847 | 659.976 | IND | IND
Y2O3, Yttrium(III) oxide (ppm) | 561.8822 | 68.7987 | 501.5192 | 622.2452 | IND | IND
Thermogravimetry
LOI, Loss On Ignition @1000°C (wt.%) | 0.6969 | 0.0711 | 0.60808 | 0.78572 | 0.65228 | 0.74152
Borate / Peroxide Fusion ICP
Al2O3, Aluminium(III) oxide (wt.%) | 8.7493 | 0.28397 | 8.5367 | 8.9619 | 8.5665 | 8.932
BaO, Barium oxide (ppm) | 1849.7328 | 77.3594 | 1802.0765 | 1897.3891 | 1800.8202 | 1898.6454
Be, Beryllium (ppm) | 15.6289 | 1.5763 | 14.1131 | 17.1446 | IND | IND
Bi, Bismuth (ppm) | 3.7965 | 0.4264 | 3.3381 | 4.2549 | IND | IND
CaO, Calcium oxide (wt.%) | 1.2537 | 0.046428 | 1.2243 | 1.2831 | 1.2244 | 1.2831
CeO2, Cerium(IV) oxide (wt.%) | 1.8762 | 0.064861 | 1.8373 | 1.9151 | 1.8291 | 1.9233
Co, Cobalt (ppm) | 18.0159 | 3.3864 | 14.7641 | 21.2677 | 15.5813 | 20.4505
Cr2O3, Chromium(III) oxide (ppm) | 577.4724 | 32.235 | 557.4767 | 597.4681 | 557.3619 | 597.5829
Cs, Cesium (ppm) | 0.27145 | 0.043588 | 0.23835 | 0.30456 | 0.20715 | 0.33575
Dy2O3, Dysprosium(III) oxide (ppm) | 204.3916 | 9.2251 | 198.9932 | 209.79 | 198.4255 | 210.3578
Er2O3, Erbium(III) oxide (ppm) | 43.6271 | 1.4904 | 42.765 | 44.4892 | 42.026 | 45.2281
Eu2O3, Europium(III) oxide (ppm) | 374.9586 | 9.6694 | 370.5735 | 379.3438 | 364.0993 | 385.818
Fe2O3, Iron(III) oxide (wt.%) | 53.2387 | 1.4398 | 52.1457 | 54.3316 | 52.0066 | 54.4708
Ga, Gallium (ppm) | 105.8017 | 21.1532 | 79.8967 | 131.7066 | 100.7282 | 110.8751
Gd2O3, Gadolinium(III) oxide (ppm) | 779.205 | 34.8161 | 757.7374 | 800.6725 | 761.8543 | 796.5556
HfO2, Hafnium dioxide (ppm) | 14.7609 | 0.90861 | 14.2327 | 15.2891 | 13.9733 | 15.5485
Ho2O3, Holmium(III) oxide (ppm) | 24.4048 | 1.5406 | 23.5078 | 25.3018 | 23.6686 | 25.1409
In, Indium (ppm) | 2.1418 | 0.16343 | 2.0031 | 2.2804 | IND | IND
K2O, Potassium oxide (wt.%) | 0.10389 | 0.01906 | 0.083124 | 0.12465 | IND | IND
La2O3, Lanthanum(III) oxide (wt.%) | 1.3729 | 0.037474 | 1.357 | 1.3888 | 1.3393 | 1.4065
Lu2O3, Lutetium(III) oxide (ppm) | 1.9171 | 0.11195 | 1.8651 | 1.9691 | 1.7883 | 2.0458
MgO, Magnesium oxide (wt.%) | 1.4021 | 0.062862 | 1.3565 | 1.4477 | 1.3813 | 1.4229
MnO, Manganese oxide (wt.%) | 0.36468 | 0.024811 | 0.34519 | 0.38417 | 0.35172 | 0.37763
Mo, Molybdenum (ppm) | 77.0732 | 4.2684 | 72.7875 | 81.3588 | 73.5556 | 80.5907
Nb2O5, Niobium(V) oxide (ppm) | 2722.8243 | 114.831 | 2636.6571 | 2808.9915 | 2612.1502 | 2833.4984
Nd2O3, Neodymium(III) oxide (wt.%) | 1.1593 | 0.036896 | 1.1402 | 1.1785 | 1.1338 | 1.1849
NiO, Nickel oxide (ppm) | 104.3283 | 20.0289 | 85.842 | 122.8147 | 85.0462 | 123.6105
P2O5, Phosphorus(V) oxide (wt.%) | 3.2271 | 0.10343 | 3.145 | 3.3092 | 3.1494 | 3.3048
PbO, Lead oxide (ppm) | 186.9196 | 13.2939 | 171.8253 | 202.0139 | 179.4741 | 194.3652
Pr6O11, Praseodymium(III,IV) oxide (ppm) | 3137.8453 | 128.0919 | 3063.1791 | 3212.5116 | 3078.6711 | 3197.0196
Rb, Rubidium (ppm) | 4.3867 | 0.46025 | 4.1715 | 4.6019 | 4.0929 | 4.6804
S, Sulphur (ppm) | 1028.3333 | 90.7092 | 940.048 | 1116.6187 | IND | IND
Sb, Antimony (ppm) | 1.8015 | 0.24593 | 1.4906 | 2.1123 | IND | IND
SiO2, Silicon dioxide (wt.%) | 20.3638 | 0.62859 | 19.6938 | 21.0337 | 19.9504 | 20.7772
Sm2O3, Samarium(III) oxide (ppm) | 1737.5013 | 43.7245 | 1718.6855 | 1756.3171 | 1677.3376 | 1797.6651
SnO2, Tin dioxide (ppm) | 58.2323 | 3.3982 | 56.2801 | 60.1845 | 55.7955 | 60.6692
SrO, Strontium oxide (ppm) | 2436.6409 | 81.7938 | 2387.4006 | 2485.8811 | 2384.4209 | 2488.8608
Ta2O5, Tantalum(V) oxide (ppm) | 29.695 | 4.2338 | 26.9148 | 32.4751 | 26.6337 | 32.7562
Tb4O7, Terbium(III,IV) oxide (ppm) | 63.2658 | 3.1728 | 61.4004 | 65.1311 | 61.4398 | 65.0917
ThO2, Thorium dioxide (ppm) | 600.1434 | 26.8156 | 584.9582 | 615.3286 | 583.9993 | 616.2875
TiO2, Titanium dioxide (wt.%) | 3.2593 | 0.099131 | 3.1842 | 3.3344 | 3.186 | 3.3326
Tm2O3, Thulium(III) oxide (ppm) | 4.0622 | 0.19936 | 3.9473 | 4.1772 | 3.9283 | 4.1961
U3O8, Uranium(V,VI) oxide (ppm) | 20.7145 | 0.49246 | 20.463 | 20.9659 | 20.039 | 21.3899
V2O5, Vanadium(V) oxide (ppm) | 412.2966 | 34.0108 | 392.7303 | 431.8629 | 395.5923 | 429.0009
WO3, Tungsten trioxide (ppm) | < 7 | IND | IND | IND | IND | IND
Y2O3, Yttrium(III) oxide (ppm) | 569.6169 | 27.0276 | 552.9545 | 586.2792 | 549.9447 | 589.2891
Yb2O3, Ytterbium(III) oxide (ppm) | 17.919 | 1.1313 | 17.3758 | 18.4621 | 17.1734 | 18.6645
ZnO, Zinc oxide (ppm) | 1533.1462 | 118.9524 | 1376.4459 | 1689.8465 | 1476.2246 | 1590.0678
ZrO2, Zirconium dioxide (ppm) | 623.7038 | 23.4756 | 605.2921 | 642.1154 | 586.1373 | 661.2702
4-Acid Digestion
Ag, Silver (ppm) | 4.431 | 0.36194 | 4.1744 | 4.6876 | 4.0907 | 4.7713
Al, Aluminium (wt.%) | 4.47 | 0.096259 | 4.4168 | 4.5232 | 4.3795 | 4.5605
As, Arsenic (ppm) | 32.9362 | 4.7153 | 29.0295 | 36.8428 | 31.0003 | 34.8721
Ba, Barium (ppm) | 1648.1311 | 71.5983 | 1609.1976 | 1687.0646 | 1617.2787 | 1678.9835
Be, Beryllium (ppm) | 15.2434 | 1.1937 | 14.4411 | 16.0457 | 14.7762 | 15.7106
Bi, Bismuth (ppm) | 4.0538 | 0.23612 | 3.8951 | 4.2126 | 3.9379 | 4.1697
Ca, Calcium (wt.%) | 0.85095 | 0.058627 | 0.81121 | 0.89068 | 0.83499 | 0.8669
Cd, Cadmium (ppm) | 0.51813 | 0.087485 | 0.4576 | 0.57866 | 0.47175 | 0.56451
Ce, Cerium (wt.%) | 1.5483 | 0.048388 | 1.5011 | 1.5955 | 1.5037 | 1.5929
Co, Cobalt (ppm) | 16.9903 | 0.93218 | 16.6155 | 17.3651 | 16.3818 | 17.5987
Cr, Chromium (ppm) | 334.6808 | 37.2057 | 302.6634 | 366.6982 | 318.7968 | 350.5647
Cs, Cesium (ppm) | 0.28775 | 0.017873 | 0.27526 | 0.30024 | 0.26636 | 0.30913
Cu, Copper (ppm) | 91.8599 | 2.1933 | 90.4906 | 93.2291 | 89.4639 | 94.2559
Dy, Dysprosium (ppm) | 181.7073 | 6.9128 | 177.336 | 186.0787 | 178.5042 | 184.9104
Er, Erbium (ppm) | 36.9362 | 1.5538 | 35.8624 | 38.01 | 36.0763 | 37.796
Eu, Europium (ppm) | 334.4206 | 18.9438 | 321.6848 | 347.1565 | 328.8897 | 339.9516
Fe, Iron (wt.%) | 35.551 | 1.5691 | 34.6214 | 36.4805 | 35.0078 | 36.0942
Gd, Gadolinium (ppm) | 670.037 | 26.3591 | 651.0753 | 688.9987 | 658.0175 | 682.0565
Hf, Hafnium (ppm) | 7.7917 | 1.3232 | 6.7379 | 8.8454 | 7.3504 | 8.2329
Ho, Holmium (ppm) | 21.3127 | 0.93131 | 20.6795 | 21.9459 | 20.8338 | 21.7916
In, Indium (ppm) | 2.1115 | 0.11819 | 2.0351 | 2.1879 | 2.0638 | 2.1592
K, Potassium (wt.%) | 0.081232 | 0.013894 | 0.072375 | 0.090089 | IND | IND
La, Lanthanum (wt.%) | 1.0644 | 0.21298 | 0.85972 | 1.2691 | 1.0328 | 1.096
Li, Lithium (ppm) | 8.2553 | 1.0667 | 7.577 | 8.9337 | 7.9074 | 8.6032
Lu, Lutetium (ppm) | 1.5722 | 0.30173 | 1.3782 | 1.7662 | 1.5095 | 1.6349
Mg, Magnesium (wt.%) | 0.82488 | 0.044558 | 0.79525 | 0.8545 | 0.80877 | 0.84098
Mn, Manganese (wt.%) | 0.26998 | 0.016316 | 0.25843 | 0.28153 | 0.26338 | 0.27658
Mo, Molybdenum (ppm) | 77.135 | 4.3974 | 74.3851 | 79.8848 | 75.4792 | 78.7908
Na, Sodium (wt.%) | 0.14862 | 0.011378 | 0.1419 | 0.15534 | IND | IND
Nd, Neodymium (wt.%) | 0.95892 | 0.025597 | 0.94212 | 0.97572 | 0.94089 | 0.97696
Ni, Nickel (ppm) | 88.5748 | 6.5484 | 84.0044 | 93.1451 | 86.0676 | 91.0819
P, Phosphorus (wt.%) | 1.3595 | 0.066673 | 1.3057 | 1.4133 | 1.3221 | 1.3969
Pb, Lead (ppm) | 204.9508 | 31.7436 | 184.2505 | 225.651 | 199.4657 | 210.4358
Pr, Praseodymium (ppm) | 2600.6204 | 86.8487 | 2515.9774 | 2685.2635 | 2523.7216 | 2677.5193
Rb, Rubidium (ppm) | 4.2208 | 0.39258 | 3.9695 | 4.4722 | 4.0175 | 4.4241
Re, Rhenium (ppm) | < 0.03 | IND | IND | IND | IND | IND
S, Sulphur (ppm) | 983.4111 | 78.8419 | 933.5973 | 1033.2249 | IND | IND
Sb, Antimony (ppm) | 1.868 | 0.1112 | 1.7986 | 1.9375 | 1.7459 | 1.9901
Sc, Scandium (ppm) | 141.3799 | 5.8076 | 137.418 | 145.3418 | 138.1328 | 144.6271
Se, Selenium (ppm) | < 50 | IND | IND | IND | IND | IND
Sm, Samarium (ppm) | 1484.8664 | 49.3775 | 1437.6604 | 1532.0725 | 1437.1586 | 1532.5743
Sn, Tin (ppm) | 35.6779 | 3.2051 | 33.5317 | 37.824 | 34.291 | 37.0647
Sr, Strontium (ppm) | 2042.8066 | 144.2724 | 1943.8925 | 2141.7206 | 2010.4634 | 2075.1497
Ta, Tantalum (ppm) | 20.9121 | 2.6948 | 18.6003 | 23.2239 | 19.7803 | 22.0439
Tb, Terbium (ppm) | 54.8891 | 2.5409 | 53.0687 | 56.7095 | 53.4438 | 56.3343
Te, Tellurium (ppm) | 0.3625 | 0.067883 | 0.30112 | 0.42387 | 0.32874 | 0.39625
Th, Thorium (ppm) | 506.3839 | 59.517 | 458.8648 | 553.903 | 496.0409 | 516.727
Ti, Titanium (wt.%) | < 1 | IND | IND | IND | IND | IND
Tl, Thallium (ppm) | 0.14535 | 0.012814 | 0.13668 | 0.15402 | IND | IND
Tm, Thulium (ppm) | 3.2956 | 0.3212 | 3.0776 | 3.5137 | 3.1935 | 3.3978
U, Uranium (ppm) | 17.6356 | 0.61435 | 17.2368 | 18.0343 | 17.1632 | 18.1079
V, Vanadium (ppm) | 207.0274 | 26.1739 | 187.7429 | 226.3118 | 202.5423 | 211.5124
W, Tungsten (ppm) | 2.3747 | 0.40124 | 2.1047 | 2.6447 | 2.2443 | 2.505
Y, Yttrium (ppm) | 448.8679 | 33.2729 | 424.5828 | 473.153 | 442.5434 | 455.1925
Yb, Ytterbium (ppm) | 14.277 | 0.93752 | 13.6036 | 14.9503 | 13.7797 | 14.7742
Zn, Zinc (ppm) | 1128.0782 | 56.7693 | 1090.3079 | 1165.8486 | 1110.3645 | 1145.792
Zr, Zirconium (ppm) | 209.7096 | 39.9025 | 179.0074 | 240.4119 | 200.3444 | 219.0749

## Indicative Values
Table 3. Indicative Values for OREAS 464
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Al2O3 | wt.% | 8.8658 | Lu2O3 | ppm | < 20 | SrO | ppm | 2366.6667
BaO | ppm | 2025 | MgO | wt.% | 1.4179 | Ta2O5 | ppm | < 100
CaO | wt.% | 1.2804 | MnO | wt.% | 0.395 | Tb4O7 | ppm | 58.1312
Cr2O3 | ppm | 475.8333 | Na2O | wt.% | 0.2345 | TiO2 | wt.% | 3.2871
Dy2O3 | ppm | 197.5 | Nb2O5 | ppm | 2772.65 | Tm2O3 | ppm | < 10
Er2O3 | ppm | 40 | NiO | ppm | 200 | U3O8 | ppm | < 100
Eu2O3 | ppm | 333.3333 | P2O5 | wt.% | 3.2545 | V2O5 | ppm | 460
Gd2O3 | ppm | 804.1667 | PbO | ppm | 300 | WO3 | ppm | < 100
HfO2 | ppm | < 100 | SiO2 | wt.% | 20.01 | Yb2O3 | ppm | 20
Ho2O3 | ppm | 20 | SnO2 | ppm | < 100 | ZnO | ppm | 1566.6667
K2O | wt.% | 0.10292 | SO3 | wt.% | 0.24817 | ZrO2 | ppm | 616.6667
Thermogravimetry
H2O- | wt.% | 0.50333
Borate / Peroxide Fusion ICP
Ag | ppm | 10.4167 | Ge | ppm | 43.3333 | Se | ppm | 40
As | ppm | 251.8708 | Li | ppm | 8.1337 | Te | ppm | < 1
B | ppm | 107.5833 | Na | wt.% | 0.13501 | Tl | ppm | < 0.5
Cd | ppm | < 1 | Re | ppm | < 0.1
Cu | ppm | 103.6968 | Sc | ppm | 151.875
4-Acid Digestion
Ga | ppm | 97.0371 | Ge | ppm | 18.3781 | Nb | ppm | 1597.7592

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 464
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
CeO2, wt.% | 1.895 | 0.027208 | 1.8406 | 1.9494 | 1.8134 | 1.9766 | 1.436% | 2.872% | 4.307% | 1.8002 | 1.9897
Fe2O3, wt.% | 54.381 | 1.031 | 52.3189 | 56.4431 | 51.2879 | 57.4741 | 1.896% | 3.792% | 5.688% | 51.6619 | 57.1001
La2O3, wt.% | 1.4064 | 0.019739 | 1.3669 | 1.4458 | 1.3471 | 1.4656 | 1.404% | 2.807% | 4.211% | 1.336 | 1.4767
Nd2O3, wt.% | 1.124 | 0.036137 | 1.0518 | 1.1963 | 1.0156 | 1.2324 | 3.215% | 6.43% | 9.645% | 1.0678 | 1.1802
Pr6O11, ppm | 3069.6676 | 98.5456 | 2872.5763 | 3266.7588 | 2774.0307 | 3365.3044 | 3.21% | 6.421% | 9.631% | 2916.1842 | 3223.151
Sm2O3, ppm | 1634.7163 | 55.2316 | 1524.253 | 1745.1795 | 1469.0214 | 1800.4111 | 3.379% | 6.757% | 10.14% | 1552.9805 | 1716.4521
ThO2, ppm | 575.7804 | 84.4666 | 406.8471 | 744.7137 | 322.3805 | 829.1803 | 14.67% | 29.34% | 44.01% | 546.9914 | 604.5694
Y2O3, ppm | 561.8822 | 68.7987 | 424.2849 | 699.4796 | 355.4862 | 768.2782 | 12.24% | 24.49% | 36.73% | 533.7881 | 589.9763
Thermogravimetry
LOI, wt.% | 0.6969 | 0.0711 | 0.5547 | 0.8391 | 0.4836 | 0.9102 | 10.2% | 20.4% | 30.61% | 0.66205 | 0.73174
Borate / Peroxide Fusion ICP
Al2O3, wt.% | 8.7493 | 0.28397 | 8.1813 | 9.3172 | 7.8973 | 9.6012 | 3.246% | 6.491% | 9.737% | 8.3118 | 9.1867
BaO, ppm | 1849.7328 | 77.3594 | 1695.014 | 2004.4516 | 1617.6546 | 2081.811 | 4.182% | 8.364% | 12.55% | 1757.2461 | 1942.2194
Be, ppm | 15.6289 | 1.5763 | 12.4763 | 18.7815 | 10.9 | 20.3578 | 10.09% | 20.17% | 30.26% | 14.8474 | 16.4103
Bi, ppm | 3.7965 | 0.4264 | 2.9437 | 4.6493 | 2.5173 | 5.0757 | 11.23% | 22.46% | 33.69% | 3.6067 | 3.9863
CaO, wt.% | 1.2537 | 0.046428 | 1.1609 | 1.3466 | 1.1145 | 1.393 | 3.703% | 7.406% | 11.11% | 1.191 | 1.3164
CeO2, wt.% | 1.8762 | 0.064861 | 1.7465 | 2.006 | 1.6817 | 2.0708 | 3.457% | 6.914% | 10.37% | 1.7824 | 1.97
Co, ppm | 18.0159 | 3.3864 | 11.2431 | 24.7887 | 7.8567 | 28.1751 | 18.8% | 37.59% | 56.39% | 17.1151 | 18.9167
Cr2O3, ppm | 577.4724 | 32.235 | 513.0024 | 641.9424 | 480.7674 | 674.1774 | 5.582% | 11.16% | 16.75% | 548.5988 | 606.346
Cs, ppm | 0.27145 | 0.043588 | 0.18428 | 0.35863 | 0.14069 | 0.40222 | 16.06% | 32.11% | 48.17% | 0.25788 | 0.28502
Dy2O3, ppm | 204.3916 | 9.2251 | 185.9415 | 222.8418 | 176.7164 | 232.0669 | 4.513% | 9.027% | 13.54% | 194.1721 | 214.6112
Er2O3, ppm | 43.6271 | 1.4904 | 40.6463 | 46.6078 | 39.156 | 48.0982 | 3.416% | 6.832% | 10.25% | 41.4457 | 45.8084
Eu2O3, ppm | 374.9586 | 9.6694 | 355.6198 | 394.2975 | 345.9504 | 403.9669 | 2.579% | 5.158% | 7.736% | 356.2107 | 393.7066
Fe2O3, wt.% | 53.2387 | 1.4398 | 50.359 | 56.1183 | 48.9192 | 57.5582 | 2.704% | 5.409% | 8.113% | 50.5767 | 55.9006
Ga, ppm | 105.8017 | 21.1532 | 63.4953 | 148.108 | 42.3421 | 169.2612 | 19.99% | 39.99% | 59.98% | 100.5116 | 111.0917
Gd2O3, ppm | 779.205 | 34.8161 | 709.5728 | 848.8371 | 674.7568 | 883.6532 | 4.468% | 8.936% | 13.4% | 740.2447 | 818.1652
HfO2, ppm | 14.7609 | 0.90861 | 12.9437 | 16.5781 | 12.0351 | 17.4867 | 6.156% | 12.31% | 18.47% | 14.0229 | 15.499
Ho2O3, ppm | 24.4048 | 1.5406 | 21.3235 | 27.486 | 19.7829 | 29.0267 | 6.313% | 12.63% | 18.94% | 23.1845 | 25.625
In, ppm | 2.1418 | 0.16343 | 1.8149 | 2.4686 | 1.6515 | 2.632 | 7.63% | 15.26% | 22.89% | 2.0347 | 2.2489
K2O, wt.% | 0.10389 | 0.01906 | 0.065769 | 0.14201 | 0.046709 | 0.16107 | 18.35% | 36.69% | 55.04% | 0.098694 | 0.10908
La2O3, wt.% | 1.3729 | 0.037474 | 1.298 | 1.4479 | 1.2605 | 1.4853 | 2.73% | 5.459% | 8.189% | 1.3043 | 1.4415
Lu2O3, ppm | 1.9171 | 0.11195 | 1.6932 | 2.141 | 1.5812 | 2.2529 | 5.839% | 11.68% | 17.52% | 1.8212 | 2.0129
MgO, wt.% | 1.4021 | 0.062862 | 1.2764 | 1.5278 | 1.2135 | 1.5907 | 4.483% | 8.967% | 13.45% | 1.332 | 1.4722
MnO, wt.% | 0.36468 | 0.024811 | 0.31505 | 0.4143 | 0.29024 | 0.43911 | 6.804% | 13.61% | 20.41% | 0.34644 | 0.38291
Mo, ppm | 77.0732 | 4.2684 | 68.5365 | 85.6099 | 64.2681 | 89.8782 | 5.538% | 11.08% | 16.61% | 73.2195 | 80.9268
Nb2O5, ppm | 2722.8243 | 114.831 | 2493.1624 | 2952.4862 | 2378.3314 | 3067.3172 | 4.217% | 8.435% | 12.65% | 2586.6831 | 2858.9655
Nd2O3, wt.% | 1.1593 | 0.036896 | 1.0856 | 1.2331 | 1.0487 | 1.27 | 3.182% | 6.365% | 9.547% | 1.1014 | 1.2173
NiO, ppm | 104.3283 | 20.0289 | 64.2705 | 144.3862 | 44.2415 | 164.4151 | 19.2% | 38.4% | 57.59% | 99.1119 | 109.5447
P2O5, wt.% | 3.2271 | 0.10343 | 3.0202 | 3.434 | 2.9168 | 3.5374 | 3.205% | 6.41% | 9.615% | 3.0657 | 3.3884
PbO, ppm | 186.9196 | 13.2939 | 160.3319 | 213.5074 | 147.038 | 226.8013 | 7.112% | 14.22% | 21.34% | 177.5736 | 196.2656
Pr6O11, ppm | 3137.8453 | 128.0919 | 2881.6615 | 3394.0292 | 2753.5696 | 3522.1211 | 4.082% | 8.164% | 12.25% | 2980.9531 | 3294.7376
Rb, ppm | 4.3867 | 0.46025 | 3.4662 | 5.3072 | 3.0059 | 5.7674 | 10.49% | 20.98% | 31.48% | 4.1673 | 4.606
S, ppm | 1028.3333 | 90.7092 | 846.9149 | 1209.7518 | 756.2057 | 1300.461 | 8.821% | 17.64% | 26.46% | 976.9167 | 1079.75
Sb, ppm | 1.8015 | 0.24593 | 1.3096 | 2.2933 | 1.0637 | 2.5392 | 13.65% | 27.3% | 40.95% | 1.7114 | 1.8915
SiO2, wt.% | 20.3638 | 0.62859 | 19.1066 | 21.621 | 18.478 | 22.2496 | 3.087% | 6.174% | 9.26% | 19.3456 | 21.382
Sm2O3, ppm | 1737.5013 | 43.7245 | 1650.0523 | 1824.9503 | 1606.3278 | 1868.6749 | 2.517% | 5.033% | 7.55% | 1650.6262 | 1824.3764
SnO2, ppm | 58.2323 | 3.3982 | 51.436 | 65.0287 | 48.0378 | 68.4268 | 5.836% | 11.67% | 17.51% | 55.3207 | 61.1439
SrO, ppm | 2436.6409 | 81.7938 | 2273.0533 | 2600.2284 | 2191.2595 | 2682.0222 | 3.357% | 6.714% | 10.07% | 2314.8088 | 2558.4729
Ta2O5, ppm | 29.695 | 4.2338 | 21.2274 | 38.1625 | 16.9936 | 42.3963 | 14.26% | 28.52% | 42.77% | 28.2102 | 31.1797
Tb4O7, ppm | 63.2658 | 3.1728 | 56.9201 | 69.6115 | 53.7472 | 72.7843 | 5.015% | 10.03% | 15.05% | 60.1025 | 66.4291
ThO2, ppm | 600.1434 | 26.8156 | 546.5122 | 653.7746 | 519.6967 | 680.5902 | 4.468% | 8.936% | 13.4% | 570.1362 | 630.1506
TiO2, wt.% | 3.2593 | 0.099131 | 3.061 | 3.4575 | 2.9619 | 3.5567 | 3.041% | 6.083% | 9.124% | 3.0963 | 3.4222
Tm2O3, ppm | 4.0622 | 0.19936 | 3.6635 | 4.4609 | 3.4641 | 4.6603 | 4.908% | 9.815% | 14.72% | 3.8591 | 4.2653
U3O8, ppm | 20.7145 | 0.49246 | 19.7296 | 21.6994 | 19.2371 | 22.1918 | 2.377% | 4.755% | 7.132% | 19.6788 | 21.7502
V2O5, ppm | 412.2966 | 34.0108 | 344.275 | 480.3183 | 310.2641 | 514.3291 | 8.249% | 16.5% | 24.75% | 391.6818 | 432.9115
WO3, ppm | < 7 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Y2O3, ppm | 569.6169 | 27.0276 | 515.5617 | 623.6721 | 488.5341 | 650.6996 | 4.745% | 9.49% | 14.23% | 541.136 | 598.0977
Yb2O3, ppm | 17.919 | 1.1313 | 15.6564 | 20.1815 | 14.5251 | 21.3128 | 6.313% | 12.63% | 18.94% | 17.023 | 18.8149
ZnO, ppm | 1533.1462 | 118.9524 | 1295.2413 | 1771.051 | 1176.2889 | 1890.0035 | 7.759% | 15.52% | 23.28% | 1456.4889 | 1609.8035
ZrO2, ppm | 623.7038 | 23.4756 | 576.7525 | 670.655 | 553.2769 | 694.1306 | 3.764% | 7.528% | 11.29% | 592.5186 | 654.8889
4-Acid Digestion
Ag, ppm | 4.431 | 0.36194 | 3.7071 | 5.1549 | 3.3452 | 5.5168 | 8.168% | 16.34% | 24.51% | 4.2094 | 4.6525
Al, wt.% | 4.47 | 0.096259 | 4.2775 | 4.6625 | 4.1812 | 4.7588 | 2.153% | 4.307% | 6.46% | 4.2465 | 4.6935
As, ppm | 32.9362 | 4.7153 | 23.5056 | 42.3668 | 18.7903 | 47.0821 | 14.32% | 28.63% | 42.95% | 31.2894 | 34.583
Ba, ppm | 1648.1311 | 71.5983 | 1504.9345 | 1791.3277 | 1433.3362 | 1862.9261 | 4.344% | 8.688% | 13.03% | 1565.7246 | 1730.5377
Be, ppm | 15.2434 | 1.1937 | 12.8559 | 17.6309 | 11.6622 | 18.8246 | 7.831% | 15.66% | 23.49% | 14.4812 | 16.0056
Bi, ppm | 4.0538 | 0.23612 | 3.5816 | 4.5261 | 3.3455 | 4.7622 | 5.825% | 11.65% | 17.47% | 3.8511 | 4.2565
Ca, wt.% | 0.85095 | 0.058627 | 0.73369 | 0.9682 | 0.67507 | 1.0268 | 6.89% | 13.78% | 20.67% | 0.8084 | 0.89349
Cd, ppm | 0.51813 | 0.087485 | 0.34316 | 0.6931 | 0.25567 | 0.78059 | 16.88% | 33.77% | 50.65% | 0.49222 | 0.54404
Ce, wt.% | 1.5483 | 0.048388 | 1.4515 | 1.6451 | 1.4031 | 1.6934 | 3.125% | 6.251% | 9.376% | 1.4709 | 1.6257
Co, ppm | 16.9903 | 0.93218 | 15.1259 | 18.8546 | 14.1937 | 19.7868 | 5.487% | 10.97% | 16.46% | 16.1408 | 17.8398
Cr, ppm | 334.6808 | 37.2057 | 260.2693 | 409.0923 | 223.0636 | 446.298 | 11.12% | 22.23% | 33.35% | 317.9468 | 351.4148
Cs, ppm | 0.28775 | 0.017873 | 0.252 | 0.3235 | 0.23413 | 0.34137 | 6.211% | 12.42% | 18.63% | 0.27336 | 0.30214
Cu, ppm | 91.8599 | 2.1933 | 87.4734 | 96.2464 | 85.2801 | 98.4397 | 2.388% | 4.775% | 7.163% | 87.2669 | 96.4529
Dy, ppm | 181.7073 | 6.9128 | 167.8817 | 195.533 | 160.9688 | 202.4458 | 3.804% | 7.609% | 11.41% | 172.622 | 190.7927
Er, ppm | 36.9362 | 1.5538 | 33.8286 | 40.0437 | 32.2749 | 41.5975 | 4.207% | 8.413% | 12.62% | 35.0894 | 38.783
Eu, ppm | 334.4206 | 18.9438 | 296.5331 | 372.3082 | 277.5893 | 391.2519 | 5.665% | 11.33% | 16.99% | 317.6996 | 351.1417
Fe, wt.% | 35.551 | 1.5691 | 32.4127 | 38.6892 | 30.8436 | 40.2583 | 4.414% | 8.827% | 13.24% | 33.7734 | 37.3285
Gd, ppm | 670.037 | 26.3591 | 617.3189 | 722.7551 | 590.9598 | 749.1142 | 3.934% | 7.868% | 11.8% | 636.5352 | 703.5389
Hf, ppm | 7.7917 | 1.3232 | 5.1454 | 10.438 | 3.8222 | 11.7611 | 16.98% | 33.96% | 50.94% | 7.4021 | 8.1813
Ho, ppm | 21.3127 | 0.93131 | 19.4501 | 23.1753 | 18.5188 | 24.1066 | 4.37% | 8.739% | 13.11% | 20.2471 | 22.3783
In, ppm | 2.1115 | 0.11819 | 1.8751 | 2.3479 | 1.757 | 2.4661 | 5.597% | 11.19% | 16.79% | 2.0059 | 2.2171
K, wt.% | 0.081232 | 0.013894 | 0.053444 | 0.10902 | 0.039551 | 0.12291 | 17.1% | 34.21% | 51.31% | 0.07717 | 0.085294
La, wt.% | 1.0644 | 0.21298 | 0.63846 | 1.4904 | 0.42548 | 1.7034 | 20.01% | 40.02% | 60.03% | 1.0112 | 1.1176
Li, ppm | 8.2553 | 1.0667 | 6.122 | 10.3887 | 5.0553 | 11.4553 | 12.92% | 25.84% | 38.76% | 7.8426 | 8.6681
Lu, ppm | 1.5722 | 0.30173 | 0.9687 | 2.1756 | 0.66696 | 2.4774 | 19.19% | 38.38% | 57.58% | 1.4936 | 1.6508
Mg, wt.% | 0.82488 | 0.044558 | 0.73576 | 0.91399 | 0.69121 | 0.95855 | 5.402% | 10.8% | 16.21% | 0.78363 | 0.86612
Mn, wt.% | 0.26998 | 0.016316 | 0.23735 | 0.30261 | 0.22103 | 0.31893 | 6.043% | 12.09% | 18.13% | 0.25648 | 0.28348
Mo, ppm | 77.135 | 4.3974 | 68.3401 | 85.9298 | 63.9427 | 90.3273 | 5.701% | 11.4% | 17.1% | 73.2782 | 80.9917
Na, wt.% | 0.14862 | 0.011378 | 0.12587 | 0.17138 | 0.11449 | 0.18275 | 7.655% | 15.31% | 22.97% | 0.14119 | 0.15605
Nd, wt.% | 0.95892 | 0.025597 | 0.90773 | 1.0101 | 0.88213 | 1.0357 | 2.669% | 5.339% | 8.008% | 0.91098 | 1.0069
Ni, ppm | 88.5748 | 6.5484 | 75.4779 | 101.6716 | 68.9294 | 108.2201 | 7.393% | 14.79% | 22.18% | 84.146 | 93.0035
P, wt.% | 1.3595 | 0.066673 | 1.2262 | 1.4929 | 1.1595 | 1.5595 | 4.904% | 9.808% | 14.71% | 1.2916 | 1.4275
Pb, ppm | 204.9508 | 31.7436 | 141.4637 | 268.4379 | 109.7201 | 300.1814 | 15.49% | 30.98% | 46.47% | 194.7032 | 215.1983
Pr, ppm | 2600.6204 | 86.8487 | 2426.923 | 2774.3179 | 2340.0742 | 2861.1666 | 3.34% | 6.679% | 10.02% | 2470.5894 | 2730.6515
Rb, ppm | 4.2208 | 0.39258 | 3.4357 | 5.006 | 3.0431 | 5.3986 | 9.301% | 18.6% | 27.9% | 4.0098 | 4.4319
Re, ppm | < 0.03 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, ppm | 983.4111 | 78.8419 | 825.7273 | 1141.0949 | 746.8854 | 1219.9369 | 8.017% | 16.03% | 24.05% | 934.2406 | 1032.5817
Sb, ppm | 1.868 | 0.1112 | 1.6456 | 2.0904 | 1.5344 | 2.2016 | 5.953% | 11.91% | 17.86% | 1.7746 | 1.9614
Sc, ppm | 141.3799 | 5.8076 | 129.7648 | 152.9951 | 123.9572 | 158.8026 | 4.108% | 8.216% | 12.32% | 134.3109 | 148.4489
Se, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sm, ppm | 1484.8664 | 49.3775 | 1386.1114 | 1583.6215 | 1336.7339 | 1632.999 | 3.325% | 6.651% | 9.976% | 1410.6231 | 1559.1097
Sn, ppm | 35.6779 | 3.2051 | 29.2676 | 42.0881 | 26.0625 | 45.2932 | 8.983% | 17.97% | 26.95% | 33.894 | 37.4617
Sr, ppm | 2042.8066 | 144.2724 | 1754.2618 | 2331.3513 | 1609.9894 | 2475.6237 | 7.062% | 14.12% | 21.19% | 1940.6662 | 2144.9469
Ta, ppm | 20.9121 | 2.6948 | 15.5225 | 26.3017 | 12.8278 | 28.9965 | 12.89% | 25.77% | 38.66% | 19.8665 | 21.9577
Tb, ppm | 54.8891 | 2.5409 | 49.8073 | 59.9708 | 47.2664 | 62.5117 | 4.629% | 9.258% | 13.89% | 52.1446 | 57.6335
Te, ppm | 0.3625 | 0.067883 | 0.22673 | 0.49826 | 0.15885 | 0.56615 | 18.73% | 37.45% | 56.18% | 0.34437 | 0.38062
Th, ppm | 506.3839 | 59.517 | 387.35 | 625.4179 | 327.833 | 684.9349 | 11.75% | 23.51% | 35.26% | 481.0647 | 531.7031
Ti, wt.% | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tl, ppm | 0.14535 | 0.012814 | 0.11972 | 0.17098 | 0.10691 | 0.18379 | 8.816% | 17.63% | 26.45% | 0.13808 | 0.15262
Tm, ppm | 3.2956 | 0.3212 | 2.6532 | 3.938 | 2.332 | 4.2592 | 9.746% | 19.49% | 29.24% | 3.1309 | 3.4604
U, ppm | 17.6356 | 0.61435 | 16.4069 | 18.8642 | 15.7925 | 19.4786 | 3.484% | 6.967% | 10.45% | 16.7538 | 18.5173
V, ppm | 207.0274 | 26.1739 | 154.6796 | 259.3751 | 128.5057 | 285.549 | 12.64% | 25.29% | 37.93% | 196.676 | 217.3787
W, ppm | 2.3747 | 0.40124 | 1.5722 | 3.1772 | 1.171 | 3.5784 | 16.9% | 33.79% | 50.69% | 2.256 | 2.4934
Y, ppm | 448.8679 | 33.2729 | 382.3221 | 515.4138 | 349.0491 | 548.6867 | 7.413% | 14.83% | 22.24% | 426.4245 | 471.3113
Yb, ppm | 14.277 | 0.93752 | 12.4019 | 16.152 | 11.4644 | 17.0895 | 6.567% | 13.13% | 19.7% | 13.5631 | 14.9908
Zn, ppm | 1128.0782 | 56.7693 | 1014.5396 | 1241.6169 | 957.7703 | 1298.3862 | 5.032% | 10.06% | 15.1% | 1071.6743 | 1184.4822
Zr, ppm | 209.7096 | 39.9025 | 129.9046 | 289.5146 | 90.0021 | 329.4171 | 19.03% | 38.05% | 57.08% | 199.2242 | 220.1951

## Fusion XRF
| Analytical results for Al2O3 in OREAS 464 (Indicative Value 8.87 wt.%) | [8 empty] | Indicative
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 8.93 | 8.72 | 8.86 | 8.97 |  |  |  | 1
| 1 | 2 | 8.91 | 8.67 | 8.83 | 9.07 |  |  |  | 1
| 1 | 3 | 8.84 | 8.75 | 8.84 | 9.02 |  |  |  | 16
| 1 | 4 | 8.84 | 8.73 | 8.9 | 9.03 |  |  |  | 8.8658
| 1 | 5 | 8.93 | 8.67 | 8.96 | 8.96 |  |  |  | 7
| 1 | 6 | 8.87 | 8.62 | 8.87 | 8.99
| Mean |  | 8.8867 | 8.6933 | 8.8767 | 9.0067
| Median |  | 8.89 | 8.695 | 8.865 | 9.005
| Std Dev. |  | 0.042269 | 0.048442 | 0.04761 | 0.041312
| Rel.Std.Dev. |  | 0.4756% | 0.5572% | 0.5363% | 0.4587%
| PDM3 |  | 0.235% | -1.946% | 0.1222% | 1.588%
| Z-Score (Absolute) |  | 0.05 | 1.95 | 0.05 | 1.3
| Analytical results for BaO in OREAS 464 (Indicative Value 2025 ppm) | [8 empty] | Indicative
BaO | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 2000 | 2000 | [5 empty] | 1
| 1 | 2 | 2200 | 2000 | [5 empty] | 2
| 1 | 3 | 2100 | 2000 | [5 empty] | 16
| 1 | 4 | 2100 | 2000 | [5 empty] | 2025
| 1 | 5 | 2000 | 1900 | [5 empty] | 8
| 1 | 6 | 1900 | 2100
| Mean |  | 2050 | 2000
| Median |  | 2050 | 2000
| Std Dev. |  | 104.8809 | 63.2456
| Rel.Std.Dev. |  | 5.116% | 3.162%
| PDM3 |  | 1.235% | -1.235%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for CaO in OREAS 464 (Indicative Value 1.28 wt.%) | [8 empty] | Indicative
CaO | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.26 | 1.29 | 1.23 | 1.33 |  |  |  | 1
| 1 | 2 | 1.27 | 1.26 | 1.26 | 1.33 |  |  |  | 3
| 1 | 3 | 1.27 | 1.27 | 1.28 | 1.32 |  |  |  | 16
| 1 | 4 | 1.26 | 1.31 | 1.23 | 1.33 |  |  |  | 1.2804
| 1 | 5 | 1.26 | 1.3 | 1.25 | 1.32 |  |  |  | 9
| 1 | 6 | 1.25 | 1.3 | 1.22 | 1.33
| Mean |  | 1.2617 | 1.2883 | 1.245 | 1.3267
| Median |  | 1.26 | 1.295 | 1.24 | 1.33
| Std Dev. |  | 0.0075277 | 0.019408 | 0.022583 | 0.005164
| Rel.Std.Dev. |  | 0.5966% | 1.506% | 1.814% | 0.3892%
| PDM3 |  | -1.464% | 0.6183% | -2.766% | 3.612%
| Z-Score (Absolute) |  | 0.41 | 0.41 | 0.93 | 1.61
| Analytical results for CeO2 in OREAS 464 (Certified Value 1.89 wt.%) | [8 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.92 | 1.92 | 1.877 | 1.87 | 2.003 | 1.904 | 1.86 | 1
| 1 | 2 | 1.94 | 1.92 | 1.856 | 1.9 | 2.003 | 1.904 | 1.84 | #N/A
| 1 | 3 | 1.93 | 1.92 | 1.856 | 1.9 | 1.993 | 1.904 | 1.85 | 16
| 1 | 4 | 1.94 | 1.9 | 1.888 | 1.88 | 1.993 | 1.904 | 1.85 | 1.895
| 1 | 5 | 1.91 | 1.92 | 1.93 | 1.89 | 1.961 | 1.904 | 1.87 | 7
| 1 | 6 | 1.92 | 1.91 | 1.888 | 1.88 | 1.993 | 1.904 | 1.86
| Mean |  | 1.9267 | 1.915 | 1.8825 | 1.8867 | 1.991 | 1.904 | 1.855
| Median |  | 1.925 | 1.92 | 1.8825 | 1.885 | 1.993 | 1.904 | 1.855
| Std Dev. |  | 0.012111 | 0.0083666 | 0.027392 | 0.012111 | 0.015492 | 0 | 0.010488
| Rel.Std.Dev. |  | 0.6286% | 0.4369% | 1.455% | 0.6419% | 0.7781% | 0% | 0.5654%
| PDM3 |  | 1.672% | 1.057% | -0.6584% | -0.4385% | 5.067% | 0.4762% | -2.11%
| Z-Score (Absolute) |  | 0.71 | 0.34 | 0.67 | 0.54 | 2.91 | 0 | 1.54
| Analytical results for Cr2O3 in OREAS 464 (Indicative Value 476 ppm) | [8 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 700 | 300 | [5 empty] | 1
| 1 | 2 | 600 | 330 | [5 empty] | 4
| 1 | 3 | 600 | 330 | [5 empty] | 16
| 1 | 4 | 600 | 300 | [5 empty] | 475.8333
| 1 | 5 | 700 | 310 | [5 empty] | 10
| 1 | 6 | 600 | 340
| Mean |  | 633.3333 | 318.3333
| Median |  | 600 | 320
| Std Dev. |  | 51.6398 | 17.224
| Rel.Std.Dev. |  | 8.154% | 5.411%
| PDM3 |  | 33.1% | -33.1%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Dy2O3 in OREAS 464 (Indicative Value 198 ppm) | [8 empty] | Indicative
Dy2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 298 | 200 | 190 |  |  |  |  | 1
| 1 | 2 | 298 | 200 | 200 |  |  |  |  | 5
| 1 | 3 | 298 | 300 | 190 |  |  |  |  | 16
| 1 | 4 | 298 | 200 | 190 |  |  |  |  | 197.5
| 1 | 5 | 298 | 200 | 200 |  |  |  |  | 11
| 1 | 6 | 298 | 200 | 200
| Mean |  | 298 | 216.6667 | 195
| Median |  | 298 | 200 | 195
| Std Dev. |  | 0 | 40.8248 | 5.4772
| Rel.Std.Dev. |  | 0% | 18.84% | 2.809%
| PDM3 |  | 50.89% | 9.705% | -1.266%
| Z-Score (Absolute) |  | 13.27 | 0 | 0.67
| Analytical results for Er2O3 in OREAS 464 (Indicative Value 40 ppm) | [8 empty] | Indicative
Er2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 100 | < 100 | 30 |  |  |  |  | 1
| 1 | 2 | < 100 | 100 | 30 |  |  |  |  | 6
| 1 | 3 | < 100 | < 100 | 30 |  |  |  |  | 16
| 1 | 4 | < 100 | < 100 | 30 |  |  |  |  | 40
| 1 | 5 | < 100 | < 100 | 30 |  |  |  |  | 12
| 1 | 6 | < 100 | < 100 | 30
| Mean |  |  | 100 | 30
| Median |  |  | 100 | 30
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  | 150% | -25%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Eu2O3 in OREAS 464 (Indicative Value 333 ppm) | [8 empty] | Indicative
Eu2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 09 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 300 | 300 | 300 |  |  |  |  | 1
| 1 | 2 | 400 | 300 | 300 |  |  |  |  | 7
| 1 | 3 | 400 | 300 | 300 |  |  |  |  | 16
| 1 | 4 | 400 | 300 | 300 |  |  |  |  | 333.3333
| 1 | 5 | 400 | 300 | 300 |  |  |  |  | 13
| 1 | 6 | 400 | 300 | 300
| Mean |  | 383.3333 | 300 | 300
| Median |  | 400 | 300 | 300
| Std Dev. |  | 40.8248 | 0 | 0
| Rel.Std.Dev. |  | 10.65% | 0% | 0%
| PDM3 |  | 15% | -10% | -10%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Fe2O3 in OREAS 464 (Certified Value 54.38 wt.%) | [8 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 09 | 13 | 14 | 15 | 22 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 54.55 | 55.5 | 53 | 53.4 | 55.1 |  |  | 1
| 1 | 2 | 54.52 | 55.5 | 53.3 | 53.3 | 55.3 |  |  | #N/A
| 1 | 3 | 54.84 | 55.9 | 53.2 | 53.3 | 55 |  |  | 16
| 1 | 4 | 54.54 | 55.9 | 53.3 | 53.2 | 55 |  |  | 54.381
| 1 | 5 | 54.74 | 55.8 | 53.5 | 53.1 | 55.3 |  |  | 8
| 1 | 6 | 54.34 | 55.7 | 53 | 53.2 | 55.1
| Mean |  | 54.5883 | 55.7167 | 53.2167 | 53.25 | 55.1333
| Median |  | 54.545 | 55.75 | 53.25 | 53.25 | 55.1
| Std Dev. |  | 0.17691 | 0.18348 | 0.19408 | 0.10488 | 0.13663
| Rel.Std.Dev. |  | 0.3241% | 0.3293% | 0.3647% | 0.197% | 0.2478%
| PDM3 |  | 0.3813% | 2.456% | -2.141% | -2.08% | 1.383%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.82 | 0.8 | 0.33
| Analytical results for Gd2O3 in OREAS 464 (Indicative Value 804 ppm) | [8 empty] | Indicative
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 09 | 14 | 22 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 800 | 800 | 700 | 800 |  |  |  | 1
| 1 | 2 | 800 | 800 | 800 | 800 |  |  |  | 8
| 1 | 3 | 800 | 900 | 800 | 800 |  |  |  | 16
| 1 | 4 | 800 | 900 | 700 | 800 |  |  |  | 804.1667
| 1 | 5 | 800 | 900 | 800 | 800 |  |  |  | 14
| 1 | 6 | 800 | 800 | 800 | 800
| Mean |  | 800 | 850 | 766.6667 | 800
| Median |  | 800 | 850 | 800 | 800
| Std Dev. |  | 0 | 54.7723 | 51.6398 | 0
| Rel.Std.Dev. |  | 0% | 6.444% | 6.736% | 0%
| PDM3 |  | -0.5181% | 5.699% | -4.663% | -0.5181%
| Z-Score (Absolute) |  | 0 | 2.02 | 1.35 | 0
| Analytical results for HfO2 in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
HfO2 | Round | Replicate | Lab | [6 empty] | 1
| No | No | 13 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | [6 empty] | 1
| 1 | 2 | < 100 | [6 empty] | 9
| 1 | 3 | < 100 | [6 empty] | 16
| 1 | 4 | < 100 | [6 empty] | < 100
| 1 | 5 | < 100 | [6 empty] | 15
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho2O3 in OREAS 464 (Indicative Value 20 ppm) | [8 empty] | Indicative
Ho2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | 20 | [5 empty] | 1
| 1 | 2 | < 100 | 20 | [5 empty] | 10
| 1 | 3 | < 100 | 20 | [5 empty] | 16
| 1 | 4 | < 100 | 20 | [5 empty] | 20
| 1 | 5 | < 100 | 20 | [5 empty] | 16
| 1 | 6 | < 100 | 20
| Mean |  |  | 20
| Median |  |  | 20
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K2O in OREAS 464 (Indicative Value 0.103 wt.%) | [8 empty] | Indicative
K2O | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.1 | 0.12 | 0.09 | 0.1 |  |  |  | 1
| 1 | 2 | 0.1 | 0.13 | 0.09 | 0.11 |  |  |  | 11
| 1 | 3 | 0.1 | 0.11 | 0.1 | 0.1 |  |  |  | 16
| 1 | 4 | 0.1 | 0.12 | 0.09 | 0.1 |  |  |  | 0.10292
| 1 | 5 | 0.1 | 0.12 | 0.09 | 0.1 |  |  |  | 17
| 1 | 6 | 0.1 | 0.12 | 0.09 | 0.1
| Mean |  | 0.1 | 0.12 | 0.091667 | 0.10167
| Median |  | 0.1 | 0.12 | 0.09 | 0.1
| Std Dev. |  | 1.5202e-17 | 0.0063246 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 1.52e-14% | 5.27% | 4.454% | 4.016%
| PDM3 |  | -2.834% | 16.6% | -10.93% | -1.215%
| Z-Score (Absolute) |  | 0.1 | 2.22 | 1.25 | 0.1
| Analytical results for La2O3 in OREAS 464 (Certified Value 1.41 wt.%) | [8 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.378 | 1.39 | 1.4 | 1.42 | 1.43 | 1.431 | 1.4 | 1
| 1 | 2 | 1.4015 | 1.39 | 1.39 | 1.39 | 1.43 | 1.431 | 1.38 | #N/A
| 1 | 3 | 1.4015 | 1.4 | 1.43 | 1.4 | 1.43 | 1.431 | 1.38 | 16
| 1 | 4 | 1.4015 | 1.38 | 1.4 | 1.42 | 1.43 | 1.443 | 1.38 | 1.4064
| 1 | 5 | 1.3898 | 1.39 | 1.38 | 1.41 | 1.43 | 1.431 | 1.4 | 9
| 1 | 6 | 1.3898 | 1.39 | 1.42 | 1.42 | 1.43 | 1.419 | 1.38
| Mean |  | 1.3937 | 1.39 | 1.4033 | 1.41 | 1.43 | 1.431 | 1.3867
| Median |  | 1.3957 | 1.39 | 1.4 | 1.415 | 1.43 | 1.431 | 1.38
| Std Dev. |  | 0.0095857 | 0.0063246 | 0.018619 | 0.012649 | 0 | 0.0075895 | 0.010328
| Rel.Std.Dev. |  | 0.6878% | 0.455% | 1.327% | 0.8971% | 0% | 0.5304% | 0.7448%
| PDM3 |  | -0.9011% | -1.163% | -0.2149% | 0.2591% | 1.681% | 1.752% | -1.4%
| Z-Score (Absolute) |  | 0.49 | 0.67 | 0 | 0.34 | 1.35 | 1.39 | 0.84
| Analytical results for Lu2O3 in OREAS 464 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Lu2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | < 20 | [5 empty] | 1
| 1 | 2 | < 100 | < 20 | [5 empty] | 1
| 1 | 3 | < 100 | < 20 | [5 empty] | 16
| 1 | 4 | < 100 | < 20 | [5 empty] | < 20
| 1 | 5 | < 100 | < 20 | [5 empty] | 7
| 1 | 6 | < 100 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for MgO in OREAS 464 (Indicative Value 1.42 wt.%) | [8 empty] | Indicative
MgO | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.45 | 1.36 | 1.39 | 1.38 |  |  |  | 1
| 1 | 2 | 1.45 | 1.43 | 1.38 | 1.39 |  |  |  | 2
| 1 | 3 | 1.46 | 1.43 | 1.42 | 1.4 |  |  |  | 16
| 1 | 4 | 1.45 | 1.44 | 1.41 | 1.38 |  |  |  | 1.4179
| 1 | 5 | 1.44 | 1.41 | 1.41 | 1.42 |  |  |  | 8
| 1 | 6 | 1.45 | 1.4 | 1.45 | 1.43
| Mean |  | 1.45 | 1.4117 | 1.41 | 1.4
| Median |  | 1.45 | 1.42 | 1.41 | 1.395
| Std Dev. |  | 0.0063246 | 0.029269 | 0.024495 | 0.020976
| Rel.Std.Dev. |  | 0.4362% | 2.073% | 1.737% | 1.498%
| PDM3 |  | 2.263% | -0.4408% | -0.5583% | -1.264%
| Z-Score (Absolute) |  | 4.53 | 0.1 | 0.1 | 1.25
| Analytical results for MnO in OREAS 464 (Indicative Value 0.395 wt.%) | [8 empty] | Indicative
MnO | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.39 | 0.42 | 0.39 | 0.39 |  |  |  | 1
| 1 | 2 | 0.39 | 0.42 | 0.38 | 0.39 |  |  |  | 3
| 1 | 3 | 0.39 | 0.42 | 0.38 | 0.38 |  |  |  | 16
| 1 | 4 | 0.39 | 0.4 | 0.38 | 0.39 |  |  |  | 0.395
| 1 | 5 | 0.39 | 0.42 | 0.38 | 0.39 |  |  |  | 9
| 1 | 6 | 0.39 | 0.42 | 0.38 | 0.39
| Mean |  | 0.39 | 0.41667 | 0.38167 | 0.38833
| Median |  | 0.39 | 0.42 | 0.38 | 0.39
| Std Dev. |  | 6.0809e-17 | 0.008165 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 1.559e-14% | 1.96% | 1.07% | 1.051%
| PDM3 |  | -1.266% | 5.485% | -3.376% | -1.688%
| Z-Score (Absolute) |  | 0.13 | 4.99 | 1.21 | 0.13
| Analytical results for Na2O in OREAS 464 (Indicative Value 0.235 wt.%) | [8 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.25 | 0.27 | 0.15 | 0.23 |  |  |  | 1
| 1 | 2 | 0.26 | 0.29 | 0.2 | 0.22 |  |  |  | 4
| 1 | 3 | 0.27 | 0.25 | 0.26 | 0.25 |  |  |  | 16
| 1 | 4 | 0.25 | 0.25 | 0.18 | 0.23 |  |  |  | 0.2345
| 1 | 5 | 0.24 | 0.27 | 0.23 | 0.23 |  |  |  | 10
| 1 | 6 | 0.25 | 0.25 | 0.14 | 0.23
| Mean |  | 0.25333 | 0.26333 | 0.19333 | 0.23167
| Median |  | 0.25 | 0.26 | 0.19 | 0.23
| Std Dev. |  | 0.010328 | 0.01633 | 0.046332 | 0.0098319
| Rel.Std.Dev. |  | 4.077% | 6.201% | 23.96% | 4.244%
| PDM3 |  | 8.031% | 12.3% | -17.56% | -1.208%
| Z-Score (Absolute) |  | 0.48 | 0.87 | 1.81 | 0.48
| Analytical results for Nb2O5 in OREAS 464 (Indicative Value 2773 ppm) | [8 empty] | Indicative
Nb2O5 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2900 | 2700 | 2720 |  |  |  |  | 1
| 1 | 2 | 2900 | 2700 | 2720 |  |  |  |  | 5
| 1 | 3 | 2900 | 2700 | 2720 |  |  |  |  | 16
| 1 | 4 | 2900 | 2700 | 2720 |  |  |  |  | 2772.65
| 1 | 5 | 2900 | 2700 | 2720 |  |  |  |  | 11
| 1 | 6 | 2900 | 2700 | 2720
| Mean |  | 2900 | 2700 | 2720
| Median |  | 2900 | 2700 | 2720
| Std Dev. |  | 0 | 0 | 0
| Rel.Std.Dev. |  | 0% | 0% | 0%
| PDM3 |  | 4.593% | -2.62% | -1.899%
| Z-Score (Absolute) |  | 6.84 | 0.67 | 0
| Analytical results for Nd2O3 in OREAS 464 (Certified Value 1.12 wt.%) | [8 empty] | Certified
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.17 | 1.13 | 1.05 | 1.1 | 1.29 | 1.143 | 1.12 | 1
| 1 | 2 | 1.18 | 1.13 | 1.05 | 1.1 | 1.3 | 1.155 | 1.1 | #N/A
| 1 | 3 | 1.16 | 1.13 | 1.08 | 1.13 | 1.3 | 1.143 | 1.1 | 16
| 1 | 4 | 1.18 | 1.13 | 1.07 | 1.14 | 1.3 | 1.155 | 1.12 | 1.124
| 1 | 5 | 1.17 | 1.13 | 1.09 | 1.11 | 1.28 | 1.155 | 1.1 | 10
| 1 | 6 | 1.18 | 1.13 | 1.05 | 1.11 | 1.29 | 1.155 | 1.12
| Mean |  | 1.1733 | 1.13 | 1.065 | 1.115 | 1.2933 | 1.151 | 1.11
| Median |  | 1.175 | 1.13 | 1.06 | 1.11 | 1.295 | 1.155 | 1.11
| Std Dev. |  | 0.008165 | 0 | 0.017607 | 0.016432 | 0.008165 | 0.0061968 | 0.010954
| Rel.Std.Dev. |  | 0.6959% | 0% | 1.653% | 1.474% | 0.6313% | 0.5384% | 0.9869%
| PDM3 |  | 4.386% | 0.5311% | -5.252% | -0.8034% | 15.06% | 2.399% | -1.248%
| Z-Score (Absolute) |  | 1.4 | 0 | 2.1 | 0.49 | 5.28 | 0.67 | 0.65
| Analytical results for NiO in OREAS 464 (Indicative Value 200 ppm) | [8 empty] | Indicative
NiO | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 200 | [6 empty] | 1
| 1 | 2 | 200 | [6 empty] | 6
| 1 | 3 | 200 | [6 empty] | 16
| 1 | 4 | 200 | [6 empty] | 200
| 1 | 5 | 200 | [6 empty] | 12
| 1 | 6 | 200
| Mean |  | 200
| Median |  | 200
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 464 (Indicative Value 3.25 wt.%) | [8 empty] | Indicative
P2O5 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 09 | 14 | 15 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 2
2
| 1 | 1 | 3.231 | 3.19 | 3.33 |  |  |  |  | 1
| 1 | 2 | 3.238 | 3.22 | 3.34 |  |  |  |  | 7
| 1 | 3 | 3.235 | 3.21 | 3.31 |  |  |  |  | 16
| 1 | 4 | 3.215 | 3.21 | 3.34 |  |  |  |  | 3.2545
| 1 | 5 | 3.232 | 3.24 | 3.3 |  |  |  |  | 13
| 1 | 6 | 3.23 | 3.18 | 3.33
| Mean |  | 3.2302 | 3.2083 | 3.325
| Median |  | 3.2315 | 3.21 | 3.33
| Std Dev. |  | 0.0079854 | 0.02137 | 0.016432
| Rel.Std.Dev. |  | 0.2472% | 0.6661% | 0.4942%
| PDM3 |  | -0.7477% | -1.419% | 2.166%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.93
| Analytical results for PbO in OREAS 464 (Indicative Value 300 ppm) | [8 empty] | Indicative
PbO | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 300 | [6 empty] | 1
| 1 | 2 | 300 | [6 empty] | 8
| 1 | 3 | 300 | [6 empty] | 16
| 1 | 4 | 300 | [6 empty] | 300
| 1 | 5 | 300 | [6 empty] | 14
| 1 | 6 | 300
| Mean |  | 300
| Median |  | 300
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr6O11 in OREAS 464 (Certified Value 3070 ppm) | [8 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 2996 | 3200 | 3720 | 3000 | 3510 | 3020 | 2900 | 1
| 1 | 2 | 3093 | 3100 | 3820 | 3000 | 3200 | 3020 | 3000 | #N/A
| 1 | 3 | 2996 | 3100 | 3410 | 3200 | 3200 | 3020 | 2900 | 16
| 1 | 4 | 2996 | 3200 | 3920 | 3100 | 2990 | 3140 | 2900 | 3069.6676
| 1 | 5 | 3093 | 3200 | 3720 | 3000 | 3200 | 3140 | 3000 | 11
| 1 | 6 | 3093 | 3100 | 3300 | 2900 | 3200 | 3140 | 3000
| Mean |  | 3044.5 | 3150 | 3648.3333 | 3033.3333 | 3216.6667 | 3080 | 2950
| Median |  | 3044.5 | 3150 | 3720 | 3000 | 3200 | 3080 | 2950
| Std Dev. |  | 53.1291 | 54.7723 | 241.5298 | 103.2796 | 166.4532 | 65.7267 | 54.7723
| Rel.Std.Dev. |  | 1.745% | 1.739% | 6.62% | 3.405% | 5.175% | 2.134% | 1.857%
| PDM3 |  | -0.8199% | 2.617% | 18.85% | -1.184% | 4.789% | 0.3366% | -3.898%
| Z-Score (Absolute) |  | 0.35 | 0.67 | 5.53 | 0.46 | 0.76 | 0 | 1.28
| Analytical results for SiO2 in OREAS 464 (Indicative Value 20.01 wt.%) | [8 empty] | Indicative
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 20.05 | 19.9 | 19.9 | 20 |  |  |  | 1
| 1 | 2 | 19.87 | 20.3 | 20 | 20 |  |  |  | 9
| 1 | 3 | 19.99 | 20.1 | 20.2 | 19.9 |  |  |  | 16
| 1 | 4 | 19.9 | 20.4 | 20 | 19.9 |  |  |  | 20.01
| 1 | 5 | 20 | 20.1 | 20.1 | 20 |  |  |  | 15
| 1 | 6 | 19.83 | 20.1 | 19.8 | 19.9
| Mean |  | 19.94 | 20.15 | 20 | 19.95
| Median |  | 19.945 | 20.1 | 20 | 19.95
| Std Dev. |  | 0.08579 | 0.17607 | 0.14142 | 0.054772
| Rel.Std.Dev. |  | 0.4302% | 0.8738% | 0.7071% | 0.2745%
| PDM3 |  | -0.3498% | 0.6997% | -0.04998% | -0.2999%
| Z-Score (Absolute) |  | 0.79 | 3.93 | 0.56 | 0.56
| Analytical results for Sm2O3 in OREAS 464 (Certified Value 1635 ppm) | [8 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 1600 | 1600 | > 5000 | 1600 | 1600 | 1860 | 1700 | 1
| 1 | 2 | 1600 | 1600 | > 5000 | 1700 | 1800 | 1860 | 1700 | #N/A
| 1 | 3 | 1600 | 1600 | > 5000 | 1600 | 1600 | 1860 | 1700 | 16
| 1 | 4 | 1600 | 1600 | > 5000 | 1600 | 1600 | 1860 | 1700 | 1634.7163
| 1 | 5 | 1600 | 1600 | > 5000 | 1600 | 1900 | 1860 | 1700 | 12
| 1 | 6 | 1600 | 1600 | > 5000 | 1700 | 1600 | 1860 | 1700
| Mean |  | 1600 | 1600 |  | 1633.3333 | 1683.3333 | 1860 | 1700
| Median |  | 1600 | 1600 |  | 1600 | 1600 | 1860 | 1700
| Std Dev. |  | 0 | 0 |  | 51.6398 | 132.916 | 0 | 0
| Rel.Std.Dev. |  | 0% | 0% |  | 3.162% | 7.896% | 0% | 0%
| PDM3 |  | -2.124% | -2.124% |  | -0.0846% | 2.974% | 13.78% | 3.994%
| Z-Score (Absolute) |  | 0.67 | 0.68 | NA | 0.06 | 0.06 | 4.04 | 1.17
| Analytical results for SnO2 in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
SnO2 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 200 | < 100 | [5 empty] | 1
| 1 | 2 | < 100 | < 100 | [5 empty] | 10
| 1 | 3 | < 100 | < 100 | [5 empty] | 16
| 1 | 4 | < 100 | < 100 | [5 empty] | < 100
| 1 | 5 | < 100 | < 100 | [5 empty] | 16
| 1 | 6 | 100 | < 100
| Mean |  | 150
| Median |  | 150
| Std Dev. |  | 70.7107
| Rel.Std.Dev. |  | 47.14%
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for SO3 in OREAS 464 (Indicative Value 0.248 wt.%) | [8 empty] | Indicative
SO3 | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.251 | [6 empty] | 1
| 1 | 2 | 0.253 | [6 empty] | 11
| 1 | 3 | 0.242 | [6 empty] | 16
| 1 | 4 | 0.235 | [6 empty] | 0.24817
| 1 | 5 | 0.256 | [6 empty] | 17
| 1 | 6 | 0.252
| Mean |  | 0.24817
| Median |  | 0.2515
| Std Dev. |  | 0.0079854
| Rel.Std.Dev. |  | 3.218%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for SrO in OREAS 464 (Indicative Value 2367 ppm) | [8 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 2500 | 2200 | [5 empty] | 1
| 1 | 2 | 2600 | 2300 | [5 empty] | 1
| 1 | 3 | 2600 | 2000 | [5 empty] | 16
| 1 | 4 | 2500 | 2300 | [5 empty] | 2366.6667
| 1 | 5 | 2600 | 2000 | [5 empty] | 7
| 1 | 6 | 2600 | 2200
| Mean |  | 2566.6667 | 2166.6667
| Median |  | 2600 | 2200
| Std Dev. |  | 51.6398 | 136.626
| Rel.Std.Dev. |  | 2.012% | 6.306%
| PDM3 |  | 8.451% | -8.451%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ta2O5 in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Ta2O5 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 500 | < 100 | [5 empty] | 1
| 1 | 2 | < 500 | < 100 | [5 empty] | 2
| 1 | 3 | < 500 | < 100 | [5 empty] | 16
| 1 | 4 | < 500 | < 100 | [5 empty] | < 100
| 1 | 5 | < 500 | < 100 | [5 empty] | 8
| 1 | 6 | < 500 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Tb4O7 in OREAS 464 (Indicative Value 58 ppm) | [8 empty] | Indicative
Tb4O7 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 106 | 60 | [5 empty] | 1
| 1 | 2 | < 106 | 60 | [5 empty] | 3
| 1 | 3 | < 106 | 80 | [5 empty] | 16
| 1 | 4 | < 106 | 60 | [5 empty] | 58.1312
| 1 | 5 | < 106 | 60 | [5 empty] | 9
| 1 | 6 | 106 | 60
| Mean |  | 106 | 63.3333
| Median |  | 106 | 60
| Std Dev. |  |  | 8.165
| Rel.Std.Dev. |  |  | 12.89%
| PDM3 |  | 82.35% | 8.949%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for ThO2 in OREAS 464 (Certified Value 576 ppm) | [8 empty] | Certified
ThO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 546 | 600 | 500 | 600 | 680 | 558 |  | 1
| 1 | 2 | 558 | 600 | 600 | 500 | 730 | 569 |  | #N/A
| 1 | 3 | 546 | 700 | 400 | 600 | 700 | 558 |  | 16
| 1 | 4 | 558 | 600 | 400 | 600 | 690 | 569 |  | 575.7804
| 1 | 5 | 546 | 600 | 400 | 600 | 650 | 558 |  | 13
| 1 | 6 | 569 | 600 | 400 | 600 | 710 | 535
| Mean |  | 553.8333 | 616.6667 | 450 | 583.3333 | 693.3333 | 557.8333
| Median |  | 552 | 600 | 400 | 600 | 695 | 558
| Std Dev. |  | 9.4745 | 40.8248 | 83.666 | 40.8248 | 27.3252 | 12.4164
| Rel.Std.Dev. |  | 1.711% | 6.62% | 18.59% | 6.999% | 3.941% | 2.226%
| PDM3 |  | -3.812% | 7.101% | -21.85% | 1.312% | 20.42% | -3.117%
| Z-Score (Absolute) |  | 0.36 | 0.99 | 2.58 | 0.28 | 2.64 | 0.28
| Analytical results for TiO2 in OREAS 464 (Indicative Value 3.29 wt.%) | [8 empty] | Indicative
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 3.24 | 3.24 | 3.29 | 3.38 |  |  |  | 1
| 1 | 2 | 3.24 | 3.23 | 3.29 | 3.36 |  |  |  | 4
| 1 | 3 | 3.25 | 3.22 | 3.3 | 3.36 |  |  |  | 16
| 1 | 4 | 3.25 | 3.26 | 3.28 | 3.38 |  |  |  | 3.2871
| 1 | 5 | 3.26 | 3.24 | 3.34 | 3.36 |  |  |  | 10
| 1 | 6 | 3.23 | 3.24 | 3.29 | 3.36
| Mean |  | 3.245 | 3.2383 | 3.2983 | 3.3667
| Median |  | 3.245 | 3.24 | 3.29 | 3.36
| Std Dev. |  | 0.010488 | 0.013292 | 0.02137 | 0.010328
| Rel.Std.Dev. |  | 0.3232% | 0.4104% | 0.6479% | 0.3068%
| PDM3 |  | -1.28% | -1.483% | 0.3422% | 2.421%
| Z-Score (Absolute) |  | 0.6 | 0.75 | 0.6 | 2.14
| Analytical results for Tm2O3 in OREAS 464 (Indicative Value < 10 ppm) | [8 empty] | Indicative
Tm2O3 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 100 | < 10 | [5 empty] | 1
| 1 | 2 | < 100 | < 10 | [5 empty] | 5
| 1 | 3 | < 100 | < 10 | [5 empty] | 16
| 1 | 4 | < 100 | < 10 | [5 empty] | < 10
| 1 | 5 | < 100 | < 10 | [5 empty] | 11
| 1 | 6 | < 100 | < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U3O8 in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
U3O8 | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 100 | < 300 | < 100 | 100 |  |  |  | 1
| 1 | 2 | < 100 | < 300 | < 100 | 100 |  |  |  | 6
| 1 | 3 | < 100 | < 300 | < 100 | < 100 |  |  |  | 16
| 1 | 4 | < 100 | < 300 | < 100 | < 100 |  |  |  | < 100
| 1 | 5 | < 100 | < 300 | < 100 | < 100 |  |  |  | 12
| 1 | 6 | < 100 | < 300 | < 100 | 100
| Mean |  |  |  |  | 100
| Median |  |  |  |  | 100
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for V2O5 in OREAS 464 (Indicative Value 460 ppm) | [8 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 500 | 420 | [5 empty] | 1
| 1 | 2 | 500 | 410 | [5 empty] | 7
| 1 | 3 | 500 | 390 | [5 empty] | 16
| 1 | 4 | 500 | 450 | [5 empty] | 460
| 1 | 5 | 500 | 430 | [5 empty] | 13
| 1 | 6 | 500 | 420
| Mean |  | 500 | 420
| Median |  | 500 | 420
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0% | 4.762%
| PDM3 |  | 8.696% | -8.696%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for WO3 in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
WO3 | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 100 | [6 empty] | 1
| 1 | 2 | < 100 | [6 empty] | 8
| 1 | 3 | < 100 | [6 empty] | 16
| 1 | 4 | < 100 | [6 empty] | < 100
| 1 | 5 | < 100 | [6 empty] | 14
| 1 | 6 | < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y2O3 in OREAS 464 (Certified Value 562 ppm) | [8 empty] | Certified
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 380 | 700 | 600 | 600 | 700 | 510 | 510 | 1
| 1 | 2 | 380 | 500 | 500 | 500 | 600 | 510 | 510 | #N/A
| 1 | 3 | 380 | 600 | 500 | 500 | 700 | 630 | 470 | 16
| 1 | 4 | 380 | 600 | 600 | 500 | 600 | 510 | 490 | 561.8822
| 1 | 5 | 380 | 600 | 600 | 600 | 700 | 510 | 490 | 14
| 1 | 6 | 380 | 600 | 600 | 600 | 600 | 510 | 510
| Mean |  | 380 | 600 | 566.6667 | 550 | 650 | 530 | 496.6667
| Median |  | 380 | 600 | 600 | 550 | 650 | 510 | 500
| Std Dev. |  | 0 | 63.2456 | 51.6398 | 54.7723 | 54.7723 | 48.9898 | 16.3299
| Rel.Std.Dev. |  | 0% | 10.54% | 9.113% | 9.959% | 8.427% | 9.243% | 3.288%
| PDM3 |  | -32.37% | 6.784% | 0.8515% | -2.115% | 15.68% | -5.674% | -11.61%
| Z-Score (Absolute) |  | 2.28 | 0.67 | 0.22 | 0 | 1.35 | 0.57 | 0.72
| Analytical results for Yb2O3 in OREAS 464 (Indicative Value 20 ppm) | [8 empty] | Indicative
Yb2O3 | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 102 | < 100 | 20 |  |  |  |  | 1
| 1 | 2 | < 102 | 100 | 20 |  |  |  |  | 9
| 1 | 3 | < 102 | 100 | 20 |  |  |  |  | 16
| 1 | 4 | < 102 | 100 | 20 |  |  |  |  | 20
| 1 | 5 | < 102 | 100 | 20 |  |  |  |  | 15
| 1 | 6 | < 102 | < 100 | 20
| Mean |  |  | 100 | 20
| Median |  |  | 100 | 20
| Std Dev. |  |  | 0 | 0
| Rel.Std.Dev. |  |  | 0% | 0%
| PDM3 |  |  | 400% | 0%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for ZnO in OREAS 464 (Indicative Value 1567 ppm) | [8 empty] | Indicative
ZnO | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 1500 | [6 empty] | 1
| 1 | 2 | 1600 | [6 empty] | 10
| 1 | 3 | 1600 | [6 empty] | 16
| 1 | 4 | 1600 | [6 empty] | 1566.6667
| 1 | 5 | 1600 | [6 empty] | 16
| 1 | 6 | 1500
| Mean |  | 1566.6667
| Median |  | 1600
| Std Dev. |  | 51.6398
| Rel.Std.Dev. |  | 3.296%
| PDM3 |  | -1.998e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for ZrO2 in OREAS 464 (Indicative Value 617 ppm) | [8 empty] | Indicative
ZrO2 | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 600 | 600 | [5 empty] | 1
| 1 | 2 | 700 | 600 | [5 empty] | 11
| 1 | 3 | 700 | 600 | [5 empty] | 16
| 1 | 4 | 700 | 500 | [5 empty] | 616.6667
| 1 | 5 | 600 | 600 | [5 empty] | 17
| 1 | 6 | 600 | 600
| Mean |  | 650 | 583.3333
| Median |  | 650 | 600
| Std Dev. |  | 54.7723 | 40.8248
| Rel.Std.Dev. |  | 8.427% | 6.999%
| PDM3 |  | 5.405% | -5.405%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 464 (Indicative Value 0.503 wt.%) | [8 empty] | Indicative
H2O- | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | at 1000°C | [6 empty] | 3
3
| 1 | 1 | 0.51 | [6 empty] | 1
| 1 | 2 | 0.49 | [6 empty] | 13
| 1 | 3 | 0.52 | [6 empty] | 16
| 1 | 4 | 0.5 | [6 empty] | 0.50333
| 1 | 5 | 0.5 | [6 empty] | 19
| 1 | 6 | 0.5
| Mean |  | 0.50333
| Median |  | 0.5
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 2.052%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI in OREAS 464 (Certified Value 0.697 wt.%) | [8 empty] | Certified
LOI | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 18 | wt.%
|  |  | at 1000°C | at 1000°C | at 1000°C | BF*XRF | at 1000°C | BF*XRF | at 1000°C | 3
3
| 1 | 1 | 0.78 | 0.64 | 0.62 | 0.8 | 1.06 | 0.632 | 1.1791 | 1
| 1 | 2 | 0.71 | 0.6 | 0.68 | 0.71 | 1.07 | 0.665 | 1.1395 | #N/A
| 1 | 3 | 0.81 | 0.63 | 0.7 | 0.81 | 1.04 | 0.639 | 1.1588 | 16
| 1 | 4 | 0.75 | 0.58 | 0.67 | 0.75 | 1.08 | 0.653 | 1.0585 | 0.6969
| 1 | 5 | 0.79 | 0.6 | 0.66 | 0.8 | 1.06 | 0.651 | 1.0598 | 16
| 1 | 6 | 0.79 | 0.62 | 0.68 | 0.75 | 1.08 | 0.679 | 1.1391
| Mean |  | 0.77167 | 0.61167 | 0.66833 | 0.77 | 1.065 | 0.65317 | 1.1225
| Median |  | 0.785 | 0.61 | 0.675 | 0.775 | 1.065 | 0.652 | 1.1393
| Std Dev. |  | 0.036009 | 0.022286 | 0.027142 | 0.039497 | 0.015166 | 0.017093 | 0.051208
| Rel.Std.Dev. |  | 4.666% | 3.643% | 4.061% | 5.129% | 1.424% | 2.617% | 4.562%
| PDM3 |  | 10.73% | -12.23% | -4.099% | 10.49% | 52.82% | -6.275% | 61.07%
| Z-Score (Absolute) |  | 0.01 | 0.91 | 0.53 | 0 | 1.7 | 0.67 | 2.03
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 464 (Indicative Value 10.4 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 20 | 22 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 5 | 12 | 5 | [12 empty] | 1
| 1 | 2 | < 5 | 18 | 5 | [12 empty] | 15
| 1 | 3 | < 5 | 18 | 5 | [12 empty] | 16
| 1 | 4 | < 5 | 12 | 5 | [12 empty] | 10.4167
| 1 | 5 | < 5 | 17 | 5 | [12 empty] | 21
| 1 | 6 | < 5 | 18 | 5
| Mean |  |  | 15.8333 | 5
| Median |  |  | 17.5 | 5
| Std Dev. |  |  | 2.9944 | 0
| Rel.Std.Dev. |  |  | 18.91% | 0%
| PDM3 |  |  | 52% | -52%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Al2O3 in OREAS 464 (Certified Value 8.75 wt.%) | [16 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 8.9789 | 8.616 | 8.484 | 8.5638 | N.A. | 8.3738 | 8.18 | 9.429 | 8.9 | 8.994 | [5 empty] | 1
| 1 | 2 | 8.9033 | 8.729 | 8.616 | 8.4443 | N.A. | 8.4776 | 8.48 | 9.126 | 8.843 | 8.956 | [5 empty] | 1
| 1 | 3 | 9.1641 | 8.692 | 8.597 | 8.5801 | N.A. | 8.4519 | 8.38 | 9.296 | 9.107 | 9.032 | [5 empty] | 16
| 1 | 4 | 8.909 | 8.673 | 8.597 | 8.3925 | N.A. | 8.5621 | 8.42 | 9.107 | 8.975 | 8.975 | [5 empty] | 8.7493
| 1 | 5 | 8.9846 | 8.616 | 8.711 | 8.3394 | N.A. | 8.5429 | 8.65 | 8.937 | 9.164 | 9.051 | [5 empty] | 18
| 1 | 6 | 8.9506 | 8.446 | 8.616 | 8.4 | N.A. | 8.5493 | 8.57 | 8.956 | 9.051 | 8.918
| Mean |  | 8.9817 | 8.6287 | 8.6035 | 8.4533 |  | 8.4929 | 8.4467 | 9.1418 | 9.0067 | 8.9877
| Median |  | 8.9648 | 8.6445 | 8.6065 | 8.4221 |  | 8.5102 | 8.45 | 9.1165 | 9.013 | 8.9845
| Std Dev. |  | 0.095588 | 0.099731 | 0.072434 | 0.097876 |  | 0.072891 | 0.16367 | 0.19193 | 0.12325 | 0.049058
| Rel.Std.Dev. |  | 1.064% | 1.156% | 0.8419% | 1.158% |  | 0.8583% | 1.938% | 2.099% | 1.368% | 0.5458%
| PDM3 |  | 2.657% | -1.378% | -1.666% | -3.382% |  | -2.93% | -3.458% | 4.487% | 2.942% | 2.725%
| Z-Score (Absolute) |  | 1.31 | 0 | 0.09 | 0.65 | NA | 0.5 | 0.67 | 1.9 | 1.4 | 1.33
| Analytical results for As in OREAS 464 (Indicative Value 252 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 42 | < 100 | 320.9894 | 330 | 20 | 570 | [9 empty] | 1
| 1 | 2 | 35 | < 100 | 331.8848 | 310 | 20 | 550 | [9 empty] | 16
| 1 | 3 | 38 | < 100 | 323.7574 | 310 | < 10 | 540 | [9 empty] | 16
| 1 | 4 | 40 | < 100 | 321.9192 | 320 | 40 | 570 | [9 empty] | 251.8708
| 1 | 5 | 41 | < 100 | 326.7707 | 310 | 40 | 550 | [9 empty] | 22
| 1 | 6 | 36 | < 100 | 323.8037 | 290 | 50 | 550
| Mean |  | 38.6667 |  | 324.8542 | 311.6667 | 34 | 555
| Median |  | 39 |  | 323.7806 | 310 | 40 | 550
| Std Dev. |  | 2.8048 |  | 3.9739 | 13.2916 | 13.4164 | 12.2474
| Rel.Std.Dev. |  | 7.254% |  | 1.223% | 4.265% | 39.46% | 2.207%
| PDM3 |  | -84.65% |  | 28.98% | 23.74% | -86.5% | 120.4%
| Z-Score (Absolute) |  | 0.76 | NA | 0.04 | 0 | 0.78 | 0.67
| Analytical results for B in OREAS 464 (Indicative Value 108 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*OES | PF*OES | [13 empty] | 0
0
| 1 | 1 | 112 | 100 | [13 empty] | 1
| 1 | 2 | 113 | 100 | [13 empty] | 17
| 1 | 3 | 112 | 120 | [13 empty] | 16
| 1 | 4 | 121 | 80 | [13 empty] | 107.5833
| 1 | 5 | 117 | 100 | [13 empty] | 23
| 1 | 6 | 116 | 100
| Mean |  | 115.1667 | 100
| Median |  | 114.5 | 100
| Std Dev. |  | 3.5449 | 12.6491
| Rel.Std.Dev. |  | 3.078% | 12.65%
| PDM3 |  | 7.049% | -7.049%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for BaO in OREAS 464 (Certified Value 1850 ppm) | [16 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1903.6 | 1736.2 | 1914.8 | 1926 | 1853.4 | 1712.9043 | 1773 | 1764.29 | 1544.3696 | N.A. | 1898 | 1812.5101 | 1819.9 | 1920 | 1518 | 1
| 1 | 2 | 1836.6 | 1758.5 | 1931.5 | 1903.6 | 1931.5 | 1715.0899 | 1819.9 | 1798.01 | 1526.2555 | N.A. | 1898 | 1851.6912 | 1842.22 | 1932 | 1485 | 2
| 1 | 3 | 1903.6 | 1775.2 | 1892.5 | 1953.9 | 1853.4 | 1760.1954 | 1787.5 | 1808.28 | 1542.8958 | N.A. | 1898 | 1801.0506 | 1898.05 | 1898 | 1664 | 16
| 1 | 4 | 1937.1 | 1909.2 | 1948.3 | 1937.1 | 1886.9 | 1680.5074 | 1817.7 | 1788.19 | 1530.4893 | N.A. | 1898 | 1819.2514 | 1752.9 | 1887 | 1518 | 1849.7328
| 1 | 5 | 1892.5 | 1764.1 | 1987.4 | 1970.6 | 1981.8 | 1720.019 | 1815.4 | 1830.39 | 1525.8 | N.A. | 1898 | 1862.0707 | 1842.22 | 1876 | 1530 | 19
| 1 | 6 | 1870.1 | 1674.8 | 1970.6 | 1948.3 | 1903.6 | 1702.7693 | 1807.6 | 1823.47 | 1543.4853 | N.A. | 1898 | 1812.0122 | 1808.73 | 1943 | 1530
| Mean |  | 1890.5833 | 1769.6667 | 1940.85 | 1939.9167 | 1901.7667 | 1715.2476 | 1803.5167 | 1802.105 | 1535.5492 |  | 1898 | 1826.431 | 1827.3367 | 1909.3333 | 1540.8333
| Median |  | 1898.05 | 1761.3 | 1939.9 | 1942.7 | 1895.25 | 1713.9971 | 1811.5 | 1803.145 | 1536.6925 |  | 1898 | 1815.8808 | 1831.06 | 1909 | 1524
| Std Dev. |  | 34.1763 | 77.1694 | 35.2455 | 23.3572 | 49.3811 | 26.1197 | 19.054 | 24.2281 | 8.9641 |  | 0 | 24.518 | 47.7195 | 26.455 | 62.5537
| Rel.Std.Dev. |  | 1.808% | 4.361% | 1.816% | 1.204% | 2.597% | 1.523% | 1.056% | 1.344% | 0.5838% |  | 0% | 1.342% | 2.611% | 1.386% | 4.06%
| PDM3 |  | 2.208% | -4.329% | 4.926% | 4.876% | 2.813% | -7.271% | -2.499% | -2.575% | -16.99% |  | 2.609% | -1.26% | -1.211% | 3.222% | -16.7%
| Z-Score (Absolute) |  | 0.55 | 0.73 | 0.98 | 0.97 | 0.64 | 0.96 | 0.2 | 0.21 | 2.5 | NA | 0.61 | 0 | 0 | 0.71 | 2.67
| Analytical results for Be in OREAS 464 (Certified Value 15.6 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 17 | 22 | 23 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 1
1
| 1 | 1 | 16.1984 | 16 | 14.4 | 10.3927 | 17 | 16 | 11 | [8 empty] | 1
| 1 | 2 | 15.5972 | 18 | 14.4 | 10.102 | 17 | 17 | 13 | [8 empty] | #N/A
| 1 | 3 | 16.009 | 17 | 14.6 | 10.1381 | 15 | 18 | 14 | [8 empty] | 16
| 1 | 4 | 16.8477 | 16 | 14.6 | 10.1341 | 16 | 15 | 13 | [8 empty] | 15.6289
| 1 | 5 | 16.701 | 18 | 15.1 | 10.2156 | 16 | 16 | 13 | [8 empty] | 20
| 1 | 6 | 16.566 | 17 | 15.7 | 10.1306 | 17 | 16 | 15
| Mean |  | 16.3199 | 17 | 14.8 | 10.1855 | 16.3333 | 16.3333 | 13.1667
| Median |  | 16.3822 | 17 | 14.6 | 10.1361 | 16.5 | 16 | 13
| Std Dev. |  | 0.47287 | 0.89443 | 0.5099 | 0.10836 | 0.8165 | 1.0328 | 1.3292
| Rel.Std.Dev. |  | 2.897% | 5.261% | 3.445% | 1.064% | 4.999% | 6.323% | 10.09%
| PDM3 |  | 4.421% | 8.773% | -5.303% | -34.83% | 4.507% | 4.507% | -15.75%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.69 | 6.08 | 0.01 | 0.01 | 3.13
| Analytical results for Bi in OREAS 464 (Certified Value 3.8 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 4.21 | 4.2 | 3.5 | 1.9 | 4 | 3.8 | [9 empty] | 1
| 1 | 2 | 4.3465 | 4.2 | 3.5 | 2.1 | 3 | 3.4 | [9 empty] | #N/A
| 1 | 3 | 4.0855 | 5.1 | 4 | 2.1 | 3 | 3.4 | [9 empty] | 16
| 1 | 4 | 3.9479 | 4.2 | 4 | 2 | 3 | 3.6 | [9 empty] | 3.7965
| 1 | 5 | 4.9774 | 4.2 | 3.5 | 2.1 | 3 | 3.5 | [9 empty] | 21
| 1 | 6 | 4.0722 | 4.2 | 4 | 2.1 | 4 | 3.7
| Mean |  | 4.2732 | 4.35 | 3.75 | 2.05 | 3.3333 | 3.5667
| Median |  | 4.1477 | 4.2 | 3.75 | 2.1 | 3 | 3.55
| Std Dev. |  | 0.37062 | 0.36742 | 0.27386 | 0.083666 | 0.5164 | 0.1633
| Rel.Std.Dev. |  | 8.673% | 8.447% | 7.303% | 4.081% | 15.49% | 4.578%
| PDM3 |  | 12.56% | 14.58% | -1.224% | -46% | -12.2% | -6.053%
| Z-Score (Absolute) |  | 0.8 | 0.91 | 0.15 | 2.71 | 0.55 | 0.15
| Analytical results for CaO in OREAS 464 (Certified Value 1.25 wt.%) | [16 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.1949 | 1.26 | 1.371 | 1.1166 | N.A. | 1.245 | 1.2092 | 1.25 | 1.353 | 1.26 | 1.26 |  |  |  |  | 1
| 1 | 2 | 1.1585 | 1.26 | 1.315 | 1.1372 | N.A. | 1.273 | 1.2032 | 1.25 | 1.2999 | 1.26 | 1.26 |  |  |  |  | 3
| 1 | 3 | 1.2201 | 1.26 | 1.287 | 1.1039 | N.A. | 1.413 | 1.2072 | 1.26 | 1.339 | 1.26 | 1.26 |  |  |  |  | 16
| 1 | 4 | 1.1529 | 1.26 | 1.315 | 1.1462 | N.A. | 1.203 | 1.1982 | 1.26 | 1.2915 | 1.26 | 1.26 |  |  |  |  | 1.2537
| 1 | 5 | 1.2355 | 1.26 | 1.301 | 1.1089 | N.A. | 1.203 | 1.2132 | 1.25 | 1.2719 | 1.26 | 1.26 |  |  |  |  | 22
| 1 | 6 | 1.2761 | 1.26 | 1.315 | 1.105 | N.A. | 1.259 | 1.2132 | 1.27 | 1.2705 | 1.26 | 1.12
| Mean |  | 1.2063 | 1.26 | 1.3173 | 1.1196 |  | 1.266 | 1.2073 | 1.2567 | 1.3043 | 1.26 | 1.2367
| Median |  | 1.2075 | 1.26 | 1.315 | 1.1127 |  | 1.252 | 1.2082 | 1.255 | 1.2957 | 1.26 | 1.26
| Std Dev. |  | 0.047289 | 0 | 0.028577 | 0.017899 |  | 0.077571 | 0.0058802 | 0.008165 | 0.0345 | 0 | 0.057155
| Rel.Std.Dev. |  | 3.92% | 0% | 2.169% | 1.599% |  | 6.127% | 0.487% | 0.6497% | 2.645% | 0% | 4.622%
| PDM3 |  | -3.781% | 0.4996% | 5.073% | -10.7% |  | 0.9781% | -3.701% | 0.2337% | 4.033% | 0.4996% | -1.362%
| Z-Score (Absolute) |  | 1.05 | 0.32 | 1.84 | 3.3 | NA | 0.26 | 1.02 | 0.26 | 1.49 | 0.32 | 0.28
| Analytical results for Cd in OREAS 464 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | < 1 | < 1 | < 10 | 0.5 | [11 empty] | 1
| 1 | 2 | < 1 | < 1 | < 10 | 0.6 | [11 empty] | 18
| 1 | 3 | < 1 | < 1 | < 10 | 0.4 | [11 empty] | 16
| 1 | 4 | < 1 | < 1 | < 10 | 0.6 | [11 empty] | < 1
| 1 | 5 | < 1 | < 1 | < 10 | 0.8 | [11 empty] | 24
| 1 | 6 | < 1 | < 1 | < 10 | 0.7
| Mean |  |  |  |  | 0.6
| Median |  |  |  |  | 0.6
| Std Dev. |  |  |  |  | 0.14142
| Rel.Std.Dev. |  |  |  |  | 23.57%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for CeO2 in OREAS 464 (Certified Value 1.88 wt.%) | [16 empty] | Certified
CeO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.9532 | 1.861 | 1.8856 | > 1.2 | 1.8057 | 1.7796 | 1.8708 | 1.8322 | 1.7178 | 1.9 | 1.9615 | 2.0758 | 1.9409 | 1.8917 | 2.0883 | 1
| 1 | 2 | 1.9224 | 1.9224 | 1.8672 | > 1.2 | 1.8242 | 1.7674 | 1.9262 | 1.8449 | 1.7816 | 1.91 | 1.9772 | 2.0752 | 1.904 | 1.8672 | 2.0514 | #N/A
| 1 | 3 | 1.9962 | 1.9777 | 1.7873 | > 1.2 | 1.8426 | 1.825 | 1.8742 | 1.8457 | 1.7645 | 1.91 | 1.9489 | 2.0364 | 1.9532 | 1.8795 | 2.0269 | 16
| 1 | 4 | 2.0023 | 1.8979 | 1.9102 | > 1.2 | 1.8426 | 1.7544 | 1.8798 | 1.8519 | 1.7474 | 1.85 | 1.9509 | 2.0394 | 1.8917 | 1.9163 | 2.076 | 1.8762
| 1 | 5 | 1.9286 | 1.8979 | 1.8426 | > 1.2 | 1.818 | 1.7914 | 1.894 | 1.872 | 1.8161 | 1.88 | 1.989 | 2.0872 | 1.9532 | 1.8549 | 2.1374 | 23
| 1 | 6 | 1.9409 | 1.9102 | 1.861 | > 1.2 | 1.9163 | 1.793 | 1.8927 | 1.8539 | 1.8175 | 1.87 | 1.9996 | 2.0842 | 1.8795 | 1.8917 | 2.1128
| Mean |  | 1.9573 | 1.9112 | 1.859 |  | 1.8416 | 1.7851 | 1.8896 | 1.8501 | 1.7741 | 1.8867 | 1.9712 | 2.0664 | 1.9204 | 1.8836 | 2.0821
| Median |  | 1.947 | 1.9041 | 1.8641 |  | 1.8334 | 1.7855 | 1.8862 | 1.8488 | 1.7731 | 1.89 | 1.9694 | 2.0755 | 1.9224 | 1.8856 | 2.0821
| Std Dev. |  | 0.034251 | 0.038529 | 0.041976 |  | 0.039321 | 0.024406 | 0.020262 | 0.013144 | 0.03918 | 0.024221 | 0.020802 | 0.022573 | 0.032655 | 0.021494 | 0.040164
| Rel.Std.Dev. |  | 1.75% | 2.016% | 2.258% |  | 2.135% | 1.367% | 1.072% | 0.7104% | 2.208% | 1.284% | 1.055% | 1.092% | 1.7% | 1.141% | 1.929%
| PDM3 |  | 4.319% | 1.863% | -0.9196% |  | -1.848% | -4.856% | 0.7136% | -1.393% | -5.441% | 0.5559% | 5.061% | 10.13% | 2.354% | 0.3898% | 10.97%
| Z-Score (Absolute) |  | 0.94 | 0.31 | 0.4 | NA | 0.83 | 1.4 | 0.02 | 0.52 | 1.54 | 0.02 | 1.12 | 2.41 | 0.44 | 0.06 | 2.63
| Analytical results for Co in OREAS 464 (Certified Value 18 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 19.8 | 20.8463 | N.A. | < 50 | 23.4654 | 5.4 | < 10 | 40 | 16.2 | [6 empty] | 1
| 1 | 2 | 19.4 | 21.1559 | N.A. | < 50 | 16.2513 | 18.7 | < 10 | 20 | 14.8 | [6 empty] | #N/A
| 1 | 3 | 19.9 | 21.1965 | N.A. | < 50 | 16.9955 | 18.7 | < 10 | 20 | 14.5 | [6 empty] | 16
| 1 | 4 | 19.4 | 20.5674 | N.A. | < 50 | 16.5757 | 4.4 | < 10 | 20 | 15.3 | [6 empty] | 18.0159
| 1 | 5 | 20.7 | 22.5026 | N.A. | < 50 | 19.2961 | 18.5 | < 10 | 40 | 15.6 | [6 empty] | 24
| 1 | 6 | 20.1 | 22.8195 | N.A. | < 50 | 17.8555 | 17.8 | < 10 | 20 | 16.4
| Mean |  | 19.8833 | 21.5147 |  |  | 18.4066 | 13.9167 |  | 26.6667 | 15.4667
| Median |  | 19.85 | 21.1762 |  |  | 17.4255 | 18.15 |  | 20 | 15.45
| Std Dev. |  | 0.48751 | 0.92223 |  |  | 2.7083 | 6.9993 |  | 10.328 | 0.75277
| Rel.Std.Dev. |  | 2.452% | 4.287% |  |  | 14.71% | 50.29% |  | 38.73% | 4.867%
| PDM3 |  | 10.37% | 19.42% |  |  | 2.169% | -22.75% |  | 48.02% | -14.15%
| Z-Score (Absolute) |  | 0.29 | 0.68 | NA | NA | 0.29 | 0.67 | NA | 1.9 | 0.75
| Analytical results for Cr2O3 in OREAS 464 (Certified Value 577 ppm) | [16 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 0
0
| 1 | 1 | 599 | 541 | 599 | 614 | 512 | < 1500 | 548.1 | 551.6037 | N.A. | 585 | 592.3678 | 290 | 1052 | 424 | 599 | 1
| 1 | 2 | 570 | 526 | 599 | 614 | 541 | < 1500 | 542.2 | 551.8551 | N.A. | 585 | 600.7397 | 440 | 1038 | 409 | 599 | 4
| 1 | 3 | 599 | 541 | 599 | 628 | 512 | < 1500 | 523.2 | 548.631 | N.A. | 585 | 589.255 | 440 | 1169 | 438 | 585 | 16
| 1 | 4 | 599 | 599 | 599 | 614 | 526 | < 1500 | 584.6 | 550.0881 | N.A. | 585 | 597.3841 | 290 | 1023 | 438 | 585 | 577.4724
| 1 | 5 | 599 | 541 | 643 | 614 | 599 | < 1500 | 565.6 | 551.1609 | N.A. | 585 | 595.8371 | 580 | 833 | 438 | 599 | 25
| 1 | 6 | 585 | 482 | 614 | 599 | 585 | < 1500 | 534.9 | 556.3916 | N.A. | 585 | 613.5295 | 580 | 1111 | 438 | 599
| Mean |  | 591.8333 | 538.3333 | 608.8333 | 613.8333 | 545.8333 |  | 549.7667 | 551.6217 |  | 585 | 598.1855 | 436.6667 | 1037.6667 | 430.8333 | 594.3333
| Median |  | 599 | 541 | 599 | 614 | 533.5 |  | 545.15 | 551.3823 |  | 585 | 596.6106 | 440 | 1045 | 438 | 599
| Std Dev. |  | 12.0734 | 37.4896 | 17.7811 | 9.1742 | 37.5948 |  | 22.14 | 2.6197 |  | 0 | 8.5055 | 129.7176 | 113.9608 | 12.0734 | 7.2296
| Rel.Std.Dev. |  | 2.04% | 6.964% | 2.921% | 1.495% | 6.888% |  | 4.027% | 0.4749% |  | 0% | 1.422% | 29.71% | 10.98% | 2.802% | 1.216%
| PDM3 |  | 2.487% | -6.778% | 5.431% | 6.297% | -5.479% |  | -4.798% | -4.477% |  | 1.304% | 3.587% | -24.38% | 79.69% | -25.39% | 2.92%
| Z-Score (Absolute) |  | 0.24 | 0.95 | 0.36 | 0.6 | 0.8 | NA | 0.71 | 0.67 | NA | 0 | 0.28 | 2.99 | 9.26 | 3.05 | 0.2
| Analytical results for Cs in OREAS 464 (Certified Value 0.27 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 2
2
| 1 | 1 | 0.29 | 0.22 | 0.31 | 0.24 | 0.29 | 0.26114 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 1
| 1 | 2 | 0.33 | 0.18 | 0.29 | 0.26 | 0.29 | 0.3387 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | #N/A
| 1 | 3 | 0.32 | 0.22 | 0.25 | 0.27 | 0.29 | 0.2876 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 16
| 1 | 4 | 0.31 | 0.28 | 0.31 | 0.24 | 0.27 | 0.23432 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 0.27145
| 1 | 5 | 0.37 | 0.21 | 0.3 | 0.26 | 0.38 | 0.27358 | 0.2 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.6 |  |  | 26
| 1 | 6 | 0.31 | 0.22 | 0.29 | 0.22 | 0.32 | 0.22366 | 0.2 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5
| Mean |  | 0.32167 | 0.22167 | 0.29167 | 0.24833 | 0.30667 | 0.26983 | 0.26667 | 0.4 |  |  |  |  | 0.51667
| Median |  | 0.315 | 0.22 | 0.295 | 0.25 | 0.29 | 0.26736 | 0.3 | 0.4 |  |  |  |  | 0.5
| Std Dev. |  | 0.027142 | 0.032506 | 0.022286 | 0.018348 | 0.039328 | 0.04133 | 0.05164 | 6.0809e-17 |  |  |  |  | 0.040825
| Rel.Std.Dev. |  | 8.438% | 14.66% | 7.641% | 7.389% | 12.82% | 15.32% | 19.36% | 1.52e-14% |  |  |  |  | 7.902%
| PDM3 |  | 18.5% | -18.34% | 7.447% | -8.517% | 12.97% | -0.5965% | -1.763% | 47.36% |  |  |  |  | 90.33%
| Z-Score (Absolute) |  | 0.67 | 1.84 | 0 | 0.97 | 0.01 | 0.49 | 0.56 | 2.44 | NA | NA | NA | NA | 4.68
| Analytical results for Cu in OREAS 464 (Indicative Value 104 ppm) | [16 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | N.A. | < 100 | 109.2261 | 75 | < 80 | 100 | 110 | [8 empty] | 1
| 1 | 2 | N.A. | < 100 | 113.029 | 364 | < 80 | 100 | 110 | [8 empty] | 19
| 1 | 3 | N.A. | < 100 | 105.7489 | 100 | < 80 | 100 | 110 | [8 empty] | 16
| 1 | 4 | N.A. | < 100 | 107.6508 | 82 | < 80 | 100 | 120 | [8 empty] | 103.6968
| 1 | 5 | N.A. | < 100 | 106.6592 | 102 | < 80 | 100 | 110 | [8 empty] | 25
| 1 | 6 | N.A. | < 100 | 108.8084 | 114 | < 80 | 100 | 110
| Mean |  |  |  | 108.5204 | 139.5 |  | 100 | 111.6667
| Median |  |  |  | 108.2296 | 101 |  | 100 | 110
| Std Dev. |  |  |  | 2.5616 | 110.8905 |  | 0 | 4.0825
| Rel.Std.Dev. |  |  |  | 2.361% | 79.49% |  | 0% | 3.656%
| PDM3 |  |  |  | 4.652% | 34.53% |  | -3.565% | 7.686%
| Z-Score (Absolute) |  | NA | NA | 0.49 | 1.12 | NA | 0.49 | 0.86
| Analytical results for Dy2O3 in OREAS 464 (Certified Value 204 ppm) | [16 empty] | Certified
Dy2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 198.55 | 203.14 | 206.59 | 206.586 | 197.4 | 215.4339 | 208.31 | 204.75 | 161.6609 | 200 | 212.9 | 183.9649 | 212.324 | 215.77 | 190.52 | 1
| 1 | 2 | 199.7 | 199.7 | 207.73 | 207.734 | 199.7 | 214.0946 | 208.77 | 202.45 | 162.5853 | 200 | 212.32 | 186.1712 | 265.119 | 218.06 | 192.81 | 5
| 1 | 3 | 200.27 | 204.29 | 206.01 | 209.455 | 209.46 | 224.0048 | 206.36 | 203.26 | 159.4608 | 200 | 208.88 | 182.28 | 265.119 | 208.88 | 191.67 | 16
| 1 | 4 | 202 | 204.86 | 220.36 | 210.603 | 200.27 | 213.8768 | 207.16 | 206.82 | 161.039 | 200 | 205.44 | 180.8659 | 207.734 | 213.47 | 195.11 | 204.3916
| 1 | 5 | 193.39 | 201.42 | 214.05 | 211.177 | 195.68 | 217.8792 | 206.59 | 212.09 | 165.0072 | 200 | 212.9 | 184.6937 | 270.857 | 208.88 | 200.85 | 27
| 1 | 6 | 199.7 | 201.42 | 211.18 | 206.012 | 205.44 | 219.9774 | 210.37 | 205.21 | 162.404 | 200 | 212.32 | 182.7004 | 262.823 | 214.62 | 199.7
| Mean |  | 198.935 | 202.4717 | 210.9867 | 208.5945 | 201.325 | 217.5445 | 207.9267 | 205.7633 | 162.0262 | 200 | 210.7933 | 183.446 | 247.3293 | 213.28 | 195.11
| Median |  | 199.7 | 202.28 | 209.455 | 208.5945 | 199.985 | 216.6566 | 207.735 | 204.98 | 162.0325 | 200 | 212.32 | 183.3327 | 263.971 | 214.045 | 193.96
| Std Dev. |  | 2.9408 | 1.9673 | 5.5133 | 2.1395 | 5.1795 | 3.9389 | 1.5269 | 3.456 | 1.8456 | 0 | 3.0284 | 1.8881 | 29.0507 | 3.731 | 4.2944
| Rel.Std.Dev. |  | 1.478% | 0.9716% | 2.613% | 1.026% | 2.573% | 1.811% | 0.7343% | 1.68% | 1.139% | 0% | 1.437% | 1.029% | 11.75% | 1.749% | 2.201%
| PDM3 |  | -2.67% | -0.9394% | 3.227% | 2.056% | -1.5% | 6.435% | 1.73% | 0.6711% | -20.73% | -2.149% | 3.132% | -10.25% | 21.01% | 4.349% | -4.541%
| Z-Score (Absolute) |  | 0.46 | 0.21 | 0.67 | 0.43 | 0.33 | 1.36 | 0.36 | 0 | 4.41 | 0.47 | 0.77 | 2.19 | 4.45 | 0.91 | 0.98
| Analytical results for Er2O3 in OREAS 464 (Certified Value 43.6 ppm) | [16 empty] | Certified
Er2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 41.28 | 41.97 | 45.51 | 44.368 | 42.77 | 43.2347 | 42.42 | 41.97 | 47.0827 | < 50 | 43.91 | 37.2726 | 26.529 | 41.17 | 44.253 | 1
| 1 | 2 | 41.17 | 43.68 | 46.77 | 44.368 | 44.48 | 44.0248 | 43.22 | 41.97 | 50.1544 | < 50 | 43.34 | 37.5851 | 32.361 | 46.88 | 42.538 | 6
| 1 | 3 | 41.97 | 44.02 | 46.08 | 44.482 | 44.6 | 43.7245 | 42.54 | 41.62 | 49.0783 | < 50 | 42.2 | 36.4822 | 32.59 | 43.45 | 44.596 | 16
| 1 | 4 | 42.77 | 44.14 | 47.68 | 45.74 | 43.57 | 43.4952 | 42.54 | 43 | 48.2762 | < 50 | 42.65 | 36.6511 | 25.729 | 44.02 | 44.368 | 43.6271
| 1 | 5 | 40.48 | 43.11 | 45.97 | 45.054 | 42.77 | 45.3041 | 42.77 | 45.51 | 52.0726 | < 50 | 43.91 | 37.174 | 33.162 | 44.02 | 50.085 | 28
| 1 | 6 | 41.74 | 42.88 | 46.31 | 45.054 | 45.17 | 42.8869 | 43.34 | 42.65 | 51.5652 | < 50 | 43.45 | 36.8426 | 32.132 | 44.02 | 43.796
| Mean |  | 41.5683 | 43.3 | 46.3867 | 44.8443 | 43.8933 | 43.7784 | 42.805 | 42.7867 | 49.7049 |  | 43.2433 | 37.0013 | 30.4172 | 43.9267 | 44.9393
| Median |  | 41.51 | 43.395 | 46.195 | 44.768 | 44.025 | 43.6098 | 42.655 | 42.31 | 49.6164 |  | 43.395 | 37.0083 | 32.2465 | 44.02 | 44.3105
| Std Dev. |  | 0.78244 | 0.81851 | 0.75654 | 0.54318 | 1.0101 | 0.84402 | 0.38687 | 1.4266 | 1.9276 |  | 0.69015 | 0.41508 | 3.3488 | 1.8207 | 2.6255
| Rel.Std.Dev. |  | 1.882% | 1.89% | 1.631% | 1.211% | 2.301% | 1.928% | 0.9038% | 3.334% | 3.878% |  | 1.596% | 1.122% | 11.01% | 4.145% | 5.842%
| PDM3 |  | -4.719% | -0.7497% | 6.325% | 2.79% | 0.6103% | 0.3468% | -1.884% | -1.926% | 13.93% |  | -0.8796% | -15.19% | -30.28% | 0.6867% | 3.008%
| Z-Score (Absolute) |  | 1.31 | 0.16 | 1.89 | 0.87 | 0.23 | 0.16 | 0.49 | 0.86 | 4.09 | NA | 0.2 | 4.34 | 8.7 | 0.26 | 0.25
| Analytical results for Eu2O3 in OREAS 464 (Certified Value 375 ppm) | [16 empty] | Certified
Eu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 375.2 | 363.58 | 376.32 | 370.53 | 379.8 | 375.819 | 398.09 | 367.17 | 233.0084 | 380 | 373.42 | 386.3653 | 274.422 | 382.11 | 365.9 | 1
| 1 | 2 | 371.7 | 361.26 | 371.69 | 376.32 | 392.5 | 367.7912 | 406.19 | 370.41 | 231.1813 | 380 | 378.05 | 385.3031 | 345.054 | 391.37 | 357.79 | 7
| 1 | 3 | 377.5 | 365.9 | 354.32 | 379.79 | 393.7 | 378.1614 | 397.39 | 357.33 | 231.5485 | 380 | 372.26 | 374.3833 | 348.528 | 379.79 | 347.37 | 16
| 1 | 4 | 383.3 | 401.79 | 380.95 | 376.32 | 384.4 | 362.6379 | 395.77 | 372.15 | 234.9471 | 370 | 373.42 | 378.1626 | 272.106 | 371.69 | 365.9 | 374.9586
| 1 | 5 | 368.2 | 376.32 | 360.11 | 376.32 | 371.7 | 371.9775 | 397.28 | 362.31 | 233.9321 | 380 | 380.37 | 385.3904 | 353.16 | 376.32 | 365.9 | 29
| 1 | 6 | 371.7 | 387.9 | 370.53 | 372.84 | 400.6 | 366.332 | 402.6 | 361.15 | 232.9413 | 380 | 379.79 | 376.9533 | 343.896 | 391.37 | 364.74
| Mean |  | 374.6 | 376.125 | 368.9867 | 375.3533 | 387.1167 | 370.4532 | 399.5533 | 365.0867 | 232.9265 | 378.3333 | 376.2183 | 381.093 | 322.861 | 382.1083 | 361.2667
| Median |  | 373.45 | 371.11 | 371.11 | 376.32 | 388.45 | 369.8843 | 397.74 | 364.74 | 232.9749 | 380 | 375.735 | 381.7328 | 344.475 | 380.95 | 365.32
| Std Dev. |  | 5.3337 | 16.0017 | 10.005 | 3.227 | 10.5139 | 5.9276 | 3.9908 | 5.7609 | 1.4174 | 4.0825 | 3.5966 | 5.191 | 38.5596 | 7.9872 | 7.5056
| Rel.Std.Dev. |  | 1.424% | 4.254% | 2.711% | 0.8597% | 2.716% | 1.6% | 0.9988% | 1.578% | 0.6085% | 1.079% | 0.956% | 1.362% | 11.94% | 2.09% | 2.078%
| PDM3 |  | -0.09565% | 0.3111% | -1.593% | 0.1053% | 3.243% | -1.202% | 6.559% | -2.633% | -37.88% | 0.9% | 0.336% | 1.636% | -13.89% | 1.907% | -3.652%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 0.67 | 0 | 1.25 | 0.52 | 2.56 | 1.09 | 15.08 | 0.32 | 0.09 | 0.61 | 5.56 | 0.72 | 1.2
| Analytical results for Fe2O3 in OREAS 464 (Certified Value 53.24 wt.%) | [16 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 50.7881 | 52.613 | 52.341 | 53.3288 | N.A. | 54.514 | 53.3388 | 32.7 | 53.485 | 54.61 | > 35.7 |  |  |  |  | 1
| 1 | 2 | 51.05 | 53.485 | 53.528 | 52.168 | N.A. | 54.314 | 53.3291 | 51.7 | 55.787 | 54.04 | > 35.7 |  |  |  |  | #N/A
| 1 | 3 | 51.7393 | 52.827 | 53.228 | 53.2758 | N.A. | 54.543 | 53.3291 | 51.3 | 55.987 | 55.33 | > 35.7 |  |  |  |  | 16
| 1 | 4 | 50.0574 | 53.414 | 52.956 | 52.7859 | N.A. | 54.357 | 53.7364 | 37.4 | 54.486 | 54.47 | > 35.7 |  |  |  |  | 53.2387
| 1 | 5 | 50.1677 | 52.999 | 53.599 | 52.9737 | N.A. | 54.414 | 53.591 | 51.9 | 54.457 | 55.47 | > 35.7 |  |  |  |  | 30
| 1 | 6 | 49.7265 | 53.885 | 53.399 | 52.8247 | N.A. | 55.015 | 53.3388 | 51.9 | 53.957 | 55.62 | > 35.7
| Mean |  | 50.5881 | 53.2038 | 53.1752 | 52.8928 |  | 54.5262 | 53.4439 | 46.15 | 54.6932 | 54.9233
| Median |  | 50.4779 | 53.2065 | 53.3135 | 52.8992 |  | 54.464 | 53.3388 | 51.5 | 54.4715 | 54.97
| Std Dev. |  | 0.74557 | 0.47334 | 0.46923 | 0.42063 |  | 0.25519 | 0.17643 | 8.7283 | 0.99707 | 0.63774
| Rel.Std.Dev. |  | 1.474% | 0.8897% | 0.8824% | 0.7952% |  | 0.468% | 0.3301% | 18.91% | 1.823% | 1.161%
| PDM3 |  | -4.979% | -0.06541% | -0.1193% | -0.6496% |  | 2.418% | 0.3855% | -13.31% | 2.732% | 3.164%
| Z-Score (Absolute) |  | 1.33 | 0 | 0.01 | 0.16 | NA | 0.67 | 0.12 | 0.77 | 0.76 | 0.88 | NA
| Analytical results for Ga in OREAS 464 (Certified Value 106 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 12 | 14 | 20 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*MS | PF*OES/MS | [6 empty] | 0
0
| 1 | 1 | 117 | 98.5 | 117.5 | 109.5 | 70.422 | N.A. | < 1 | 184 | 68 | [6 empty] | 1
| 1 | 2 | 118 | 102.5 | 115 | 109.5 | 71.1984 | N.A. | < 1 | 184 | 69 | [6 empty] | 21
| 1 | 3 | 117 | 100.5 | 119.5 | 107.5 | 70.6778 | N.A. | < 1 | 178 | 62 | [6 empty] | 16
| 1 | 4 | 125 | 104 | 119 | 105.5 | 70.6153 | N.A. | < 1 | 172 | 64 | [6 empty] | 105.8017
| 1 | 5 | 116.5 | 105.5 | 119 | 150.5 | 71.4563 | N.A. | < 1 | 174 | 73 | [6 empty] | 31
| 1 | 6 | 118.5 | 102 | 113.5 | 145.5 | 70.7797 | N.A. | < 1 | 180 | 69
| Mean |  | 118.6667 | 102.1667 | 117.25 | 121.3333 | 70.8583 |  |  | 178.6667 | 67.5
| Median |  | 117.5 | 102.25 | 118.25 | 109.5 | 70.7287 |  |  | 179 | 68.5
| Std Dev. |  | 3.1885 | 2.4833 | 2.4648 | 20.7694 | 0.39026 |  |  | 5.0067 | 3.937
| Rel.Std.Dev. |  | 2.687% | 2.431% | 2.102% | 17.12% | 0.5508% |  |  | 2.802% | 5.833%
| PDM3 |  | 12.16% | -3.436% | 10.82% | 14.68% | -33.03% |  |  | 68.87% | -36.2%
| Z-Score (Absolute) |  | 0.01 | 0.67 | 0 | 0.18 | 2.07 | NA | NA | 2.75 | 2.22
| Analytical results for Gd2O3 in OREAS 464 (Certified Value 779 ppm) | [16 empty] | Certified
Gd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 783.77 | 784.9 | 764.17 | 715.76 | 795.29 | 767.29 | 722.15 | 821.6378 | 830 | 805.09 | 591 | 901.333 | 781.5 | 735.36 |  | 1
| 1 | 2 | 786.07 | 776.9 | 775.7 | 730.75 | 839.09 | 783.54 | 743.07 | 819.6602 | 830 | 806.82 | 611 | 967.031 | 818.3 | 719.22 |  | 8
| 1 | 3 | 783.77 | 794.1 | 764.17 | 733.05 | 783.77 | 768.09 | 723.257 | 822.7532 | 820 | 796.45 | 616 | 986.626 | 783.8 | 703.09 |  | 16
| 1 | 4 | 825.26 | 797.6 | 795.29 | 737.66 | 821.8 | 772.82 | 728.547 | 820.9673 | 800 | 798.75 | 610 | 992.389 | 797.6 | 729.6 |  | 779.205
| 1 | 5 | 801.06 | 791.8 | 774.55 | 731.9 | 801.06 | 771.78 | 725.665 | 818.8407 | 820 | 812.58 | 625 | 988.931 | 781.5 | 751.5 |  | 32
| 1 | 6 | 781.46 | 773.4 | 776.85 | 724.99 | 759.56 | 773.86 | 726.357 | 824.4843 | 820 | 820.65 | 623 | 948.59 | 797.6 | 741.12
| Mean |  | 793.565 | 786.45 | 775.1217 | 729.0183 | 800.095 | 772.8967 | 728.1743 | 821.3906 | 820 | 806.7233 | 612.6667 | 964.15 | 793.3833 | 729.9817
| Median |  | 784.92 | 788.35 | 775.125 | 731.325 | 798.175 | 772.3 | 726.011 | 821.3025 | 820 | 805.955 | 613.5 | 976.8285 | 790.7 | 732.48
| Std Dev. |  | 17.0607 | 9.7506 | 11.4043 | 7.6713 | 28.019 | 5.831 | 7.6432 | 2.0574 | 10.9545 | 8.9448 | 12.242 | 34.966 | 14.3574 | 17.0596
| Rel.Std.Dev. |  | 2.15% | 1.24% | 1.471% | 1.052% | 3.502% | 0.7544% | 1.05% | 0.2505% | 1.336% | 1.109% | 1.998% | 3.627% | 1.81% | 2.337%
| PDM3 |  | 1.843% | 0.9298% | -0.524% | -6.441% | 2.681% | -0.8096% | -6.549% | 5.414% | 5.235% | 3.532% | -21.37% | 23.74% | 1.82% | -6.317%
| Z-Score (Absolute) |  | 0.01 | 0.01 | 0.3 | 1.47 | 0.34 | 0.35 | 1.49 | 0.88 | 0.84 | 0.51 | 4.43 | 4.51 | 0.17 | 1.45
| Analytical results for Ge in OREAS 464 (Indicative Value 43.3 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 20 | 22 | 23 | [12 empty] | ppm
|  |  | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 1
1
| 1 | 1 | < 1 | 40 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | 60 | < 1 | [12 empty] | 15
| 1 | 3 | < 1 | 40 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | 40 | < 1 | [12 empty] | 43.3333
| 1 | 5 | < 1 | 40 | < 1 | [12 empty] | 21
| 1 | 6 | < 1 | 40 | < 1
| Mean |  |  | 43.3333
| Median |  |  | 40
| Std Dev. |  |  | 8.165
| Rel.Std.Dev. |  |  | 18.84%
| PDM3 |  |  | 8.882e-14%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for HfO2 in OREAS 464 (Certified Value 14.8 ppm) | [16 empty] | Certified
HfO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 20 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 15.3 | 15.33 | 15.3 | 15.21 | 14.98 | 15.68 | 12.03 | N.A. | 14.2 | 13 | 14.2 |  |  |  |  | 1
| 1 | 2 | 15.3 | 14.51 | 15.3 | 14.98 | 15.33 | 15.1 | 11.79 | N.A. | 14.2 | 15.3 | 13 |  |  |  |  | 9
| 1 | 3 | 15.3 | 15.68 | 16.5 | 14.86 | 14.39 | 14.86 | 11.44 | N.A. | 13 | 15.3 | 14.2 |  |  |  |  | 16
| 1 | 4 | 15.3 | 17.1 | 16.5 | 15.33 | 14.51 | 14.74 | 11.44 | N.A. | 14.2 | 13 | 14.2 |  |  |  |  | 14.7609
| 1 | 5 | 14.2 | 14.98 | 16.5 | 15.33 | 14.27 | 14.86 | 11.68 | N.A. | 14.2 | 15.3 | 14.2 |  |  |  |  | 33
| 1 | 6 | 15.3 | 14.51 | 15.3 | 14.27 | 13.68 | 14.74 | 11.79 | N.A. | 14.2 | 14.2 | 13
| Mean |  | 15.1167 | 15.3517 | 15.9 | 14.9967 | 14.5267 | 14.9967 | 11.695 |  | 14 | 14.35 | 13.8
| Median |  | 15.3 | 15.155 | 15.9 | 15.095 | 14.45 | 14.86 | 11.735 |  | 14.2 | 14.75 | 14.2
| Std Dev. |  | 0.44907 | 0.97158 | 0.65727 | 0.40367 | 0.5749 | 0.35965 | 0.22828 |  | 0.4899 | 1.1292 | 0.61968
| Rel.Std.Dev. |  | 2.971% | 6.329% | 4.134% | 2.692% | 3.958% | 2.398% | 1.952% |  | 3.499% | 7.869% | 4.49%
| PDM3 |  | 2.41% | 4.002% | 7.717% | 1.597% | -1.587% | 1.597% | -20.77% |  | -5.155% | -2.784% | -6.51%
| Z-Score (Absolute) |  | 0.54 | 0.81 | 1.51 | 0.37 | 0.2 | 0.2 | 3.67 | NA | 0.9 | 0.42 | 1.14
| Analytical results for Ho2O3 in OREAS 464 (Certified Value 24.4 ppm) | [16 empty] | Certified
Ho2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 23.02 | 24.056 | 25.888 | 24.628 | 23.83 | 25.8856 | 24.86 | 23.37 | 17.0128 | < 50 | 24.63 | 21.753 | 24.972 | 26.12 | 22.223 | 1
| 1 | 2 | 23.25 | 23.597 | 25.659 | 25.201 | 24.28 | 25.6636 | 24.97 | 23.71 | 16.0863 | < 50 | 24.28 | 21.8912 | 26.461 | 26.8 | 21.879 | 10
| 1 | 3 | 23.02 | 24.399 | 25.659 | 24.743 | 25.32 | 26.2851 | 25.32 | 23.94 | 16.5143 | < 50 | 24.06 | 21.2398 | 27.034 | 25.43 | 21.421 | 16
| 1 | 4 | 23.83 | 26.576 | 26.919 | 25.201 | 24.17 | 25.1175 | 24.86 | 23.94 | 16.2422 | < 50 | 24.06 | 21.1782 | 24.514 | 25.43 | 21.879 | 24.4048
| 1 | 5 | 22.57 | 24.857 | 26.003 | 25.201 | 23.25 | 26.1269 | 25.2 | 25.09 | 16.5053 | < 50 | 24.4 | 21.8063 | 27.034 | 25.43 | 22.223 | 34
| 1 | 6 | 23.14 | 23.254 | 26.576 | 25.086 | 24.86 | 25.4647 | 25.09 | 23.71 | 16.1865 | < 50 | 24.74 | 21.2866 | 25.774 | 26.12 | 22.223
| Mean |  | 23.1383 | 24.4565 | 26.1173 | 25.01 | 24.285 | 25.7572 | 25.05 | 23.96 | 16.4246 |  | 24.3617 | 21.5259 | 25.9648 | 25.8883 | 21.9747
| Median |  | 23.08 | 24.2275 | 25.9455 | 25.1435 | 24.225 | 25.7746 | 25.03 | 23.825 | 16.3737 |  | 24.34 | 21.5198 | 26.1175 | 25.775 | 22.051
| Std Dev. |  | 0.41034 | 1.1833 | 0.51746 | 0.25785 | 0.73394 | 0.43254 | 0.1874 | 0.59174 | 0.33613 |  | 0.2847 | 0.32362 | 1.0639 | 0.56012 | 0.31933
| Rel.Std.Dev. |  | 1.773% | 4.838% | 1.981% | 1.031% | 3.022% | 1.679% | 0.7481% | 2.47% | 2.046% |  | 1.169% | 1.503% | 4.097% | 2.164% | 1.453%
| PDM3 |  | -5.189% | 0.2119% | 7.017% | 2.48% | -0.4908% | 5.542% | 2.644% | -1.823% | -32.7% |  | -0.1767% | -11.8% | 6.392% | 6.079% | -9.958%
| Z-Score (Absolute) |  | 0.65 | 0.02 | 0.88 | 0.31 | 0.06 | 0.69 | 0.33 | 0.35 | 4.11 | NA | 0.02 | 1.48 | 0.8 | 0.76 | 1.25
| Analytical results for In in OREAS 464 (Certified Value 2.14 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 2.0433 | 2.1 | 2 | 1.5 | 1.8 | 2.3 | [9 empty] | 1
| 1 | 2 | 2.1468 | 2.2 | 2 | 1.7 | 2.2 | 2.2 | [9 empty] | #N/A
| 1 | 3 | 2.5078 | 2.1 | 2 | 1.7 | 2 | 2.1 | [9 empty] | 16
| 1 | 4 | 2.2575 | 2.1 | 2 | 1.5 | 2 | 2.6 | [9 empty] | 2.1418
| 1 | 5 | 2.2685 | 2.3 | 2 | 1.7 | 2.2 | 2.2 | [9 empty] | 35
| 1 | 6 | 2.0491 | 2.1 | 2 | 1.6 | 2.2 | 3.5
| Mean |  | 2.2122 | 2.15 | 2 | 1.6167 | 2.0667 | 2.4833
| Median |  | 2.2022 | 2.1 | 2 | 1.65 | 2.1 | 2.25
| Std Dev. |  | 0.17438 | 0.083666 | 0 | 0.098319 | 0.1633 | 0.52694
| Rel.Std.Dev. |  | 7.883% | 3.891% | 0% | 6.082% | 7.902% | 21.22%
| PDM3 |  | 3.287% | 0.3844% | -6.619% | -24.52% | -3.506% | 15.95%
| Z-Score (Absolute) |  | 0.66 | 0.26 | 0.69 | 3.13 | 0.26 | 1.09
| Analytical results for K2O in OREAS 464 (Certified Value 0.104 wt.%) | [16 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.265 | 0.145 | 0.096 | 0.11809 | N.A. | < 0.12 | 0.0861 | 0.1 | 0.096 | < 0.12 | [5 empty] | 1
| 1 | 2 | 0.2614 | 0.145 | 0.084 | 0.11796 | N.A. | < 0.12 | 0.0841 | 0.1 | 0.096 | < 0.12 | [5 empty] | 11
| 1 | 3 | 0.2734 | 0.133 | 0.084 | 0.11879 | N.A. | < 0.12 | 0.0861 | 0.1 | 0.096 | < 0.12 | [5 empty] | 16
| 1 | 4 | 0.2542 | 0.133 | 0.096 | 0.11892 | N.A. | < 0.12 | 0.0851 | 0.11 | 0.084 | < 0.12 | [5 empty] | 0.10389
| 1 | 5 | 0.2686 | 0.133 | 0.096 | 0.11747 | N.A. | < 0.12 | 0.0883 | 0.1 | 0.084 | < 0.12 | [5 empty] | 36
| 1 | 6 | 0.3011 | 0.133 | 0.096 | 0.11768 | N.A. | < 0.12 | 0.0881 | 0.1 | 0.072 | < 0.12
| Mean |  | 0.27062 | 0.137 | 0.092 | 0.11815 |  |  | 0.0863 | 0.10167 | 0.088
| Median |  | 0.2668 | 0.133 | 0.096 | 0.11803 |  |  | 0.0861 | 0.1 | 0.09
| Std Dev. |  | 0.016291 | 0.0061968 | 0.0061968 | 0.00058883 |  |  | 0.0016492 | 0.0040825 | 0.009798
| Rel.Std.Dev. |  | 6.02% | 4.523% | 6.736% | 0.4984% |  |  | 1.911% | 4.016% | 11.13%
| PDM3 |  | 160.5% | 31.87% | -11.44% | 13.73% |  |  | -16.93% | -2.138% | -15.29%
| Z-Score (Absolute) |  | 7.15 | 1.53 | 0.41 | 0.72 | NA | NA | 0.67 | 0 | 0.58 | NA
| Analytical results for La2O3 in OREAS 464 (Certified Value 1.37 wt.%) | [16 empty] | Certified
La2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.3898 | 1.3253 | 1.3722 | > 1.2 | 1.3663 | 1.2631 | 1.3741 | 1.3726 | 1.3691 | 1.35 | 1.4057 | 1.4782 | 1.2901 | 1.3956 | 1.3487 | 1
| 1 | 2 | 1.3722 | 1.3898 | 1.378 | > 1.2 | 1.378 | 1.2549 | 1.4171 | 1.3833 | 1.3008 | 1.36 | 1.4065 | 1.4781 | 1.3487 | 1.4074 | 1.2784 | #N/A
| 1 | 3 | 1.425 | 1.3839 | 1.3077 | > 1.2 | 1.4015 | 1.2985 | 1.3981 | 1.3827 | 1.2833 | 1.36 | 1.3977 | 1.4382 | 1.3956 | 1.4074 | 1.3135 | 16
| 1 | 4 | 1.425 | 1.3253 | 1.4132 | > 1.2 | 1.3839 | 1.2351 | 1.3902 | 1.3783 | 1.3751 | 1.33 | 1.398 | 1.462 | 1.337 | 1.4191 | 1.466 | 1.3729
| 1 | 5 | 1.378 | 1.337 | 1.3487 | > 1.2 | 1.3722 | 1.2921 | 1.4008 | 1.3884 | 1.3411 | 1.35 | 1.4262 | 1.4989 | 1.3839 | 1.3722 | 1.4191 | 37
| 1 | 6 | 1.378 | 1.3194 | 1.3722 | > 1.2 | 1.4425 | 1.2476 | 1.3876 | 1.3845 | 1.3444 | 1.34 | 1.4429 | 1.4787 | 1.337 | 1.4074 | 1.3604
| Mean |  | 1.3947 | 1.3468 | 1.3653 |  | 1.3907 | 1.2652 | 1.3947 | 1.3816 | 1.3356 | 1.3483 | 1.4129 | 1.4724 | 1.3487 | 1.4015 | 1.3643
| Median |  | 1.3839 | 1.3312 | 1.3722 |  | 1.3809 | 1.259 | 1.3942 | 1.383 | 1.3427 | 1.35 | 1.4061 | 1.4782 | 1.3429 | 1.4074 | 1.3545
| Std Dev. |  | 0.024184 | 0.031612 | 0.03507 |  | 0.028089 | 0.025105 | 0.014457 | 0.0054675 | 0.036717 | 0.01169 | 0.018013 | 0.020438 | 0.037822 | 0.016171 | 0.06865
| Rel.Std.Dev. |  | 1.734% | 2.347% | 2.569% |  | 2.02% | 1.984% | 1.037% | 0.3957% | 2.749% | 0.867% | 1.275% | 1.388% | 2.804% | 1.154% | 5.032%
| PDM3 |  | 1.585% | -1.902% | -0.5512% |  | 1.299% | -7.844% | 1.585% | 0.6362% | -2.715% | -1.79% | 2.91% | 7.244% | -1.761% | 2.084% | -0.6229%
| Z-Score (Absolute) |  | 0.6 | 0.72 | 0.21 | NA | 0.21 | 2.98 | 0.6 | 0.24 | 1.03 | 0.68 | 1.11 | 2.76 | 0.67 | 0.79 | 0.24
| Analytical results for Li in OREAS 464 (Indicative Value 8.13 ppm) | [16 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 8.8475 | 8 | 9 | < 10 | 7 | < 10 | [9 empty] | 1
| 1 | 2 | 8.8174 | 8 | 7 | < 10 | 7 | < 10 | [9 empty] | 16
| 1 | 3 | 8.1222 | 8 | 8 | < 10 | 8 | < 10 | [9 empty] | 16
| 1 | 4 | 8.8315 | 8 | 8 | < 10 | 7 | < 10 | [9 empty] | 8.1337
| 1 | 5 | 9.1703 | 8 | 8 | < 10 | 8 | < 10 | [9 empty] | 22
| 1 | 6 | 8.6747 | 8 | 9 | < 10 | 8 | < 10
| Mean |  | 8.7439 | 8 | 8.1667 |  | 7.5
| Median |  | 8.8245 | 8 | 8 |  | 7.5
| Std Dev. |  | 0.34554 | 0 | 0.75277 |  | 0.54772
| Rel.Std.Dev. |  | 3.952% | 0% | 9.218% |  | 7.303%
| PDM3 |  | 7.502% | -1.644% | 0.4048% |  | -7.791%
| Z-Score (Absolute) |  | 1.59 | 0.17 | 0.17 | NA | 1.18 | NA
| Analytical results for Lu2O3 in OREAS 464 (Certified Value 1.92 ppm) | [16 empty] | Certified
Lu2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 2.047 | 1.899 | 1.831 | 2.001 | 1.944 | 2.0333 | 1.82 | 1.93 | < 50 | 2.27 | 1.6472 | 1.012 | 1.82 | 2.183 |  | 1
| 1 | 2 | 1.967 | 1.922 | 1.979 | 2.024 | 1.956 | 2.1948 | 1.71 | 1.71 | < 50 | 2.27 | 1.6453 | 1.046 | 1.82 | 2.217 |  | 1
| 1 | 3 | 1.99 | 1.797 | 1.842 | 2.001 | 1.842 | 1.9303 | 1.93 | 1.82 | < 50 | 2.27 | 1.628 | 1.08 | 1.82 | 2.229 |  | 16
| 1 | 4 | 2.035 | 1.865 | 1.99 | 2.024 | 1.967 | 1.7098 | 1.82 | 1.82 | < 50 | 2.27 | 1.5808 | 0.876 | 2.05 | 2.274 |  | 1.9171
| 1 | 5 | 1.899 | 1.888 | 1.808 | 2.092 | 1.979 | 2.079 | 1.93 | 2.16 | < 50 | 2.27 | 1.6445 | 1.012 | 1.82 | 2.502 |  | 38
| 1 | 6 | 1.922 | 1.831 | 1.956 | 2.126 | 1.876 | 1.8459 | 1.82 | 1.82 | < 50 | 1.71 | 1.589 | 1.046 | 1.82 | 2.183
| Mean |  | 1.9767 | 1.867 | 1.901 | 2.0447 | 1.9273 | 1.9655 | 1.8383 | 1.8767 |  | 2.1767 | 1.6225 | 1.012 | 1.8583 | 2.2647
| Median |  | 1.9785 | 1.8765 | 1.899 | 2.024 | 1.95 | 1.9818 | 1.82 | 1.82 |  | 2.27 | 1.6363 | 1.029 | 1.82 | 2.223
| Std Dev. |  | 0.059399 | 0.046239 | 0.082535 | 0.052006 | 0.055244 | 0.17378 | 0.082805 | 0.15526 |  | 0.22862 | 0.03004 | 0.071319 | 0.093897 | 0.12108
| Rel.Std.Dev. |  | 3.005% | 2.477% | 4.342% | 2.544% | 2.866% | 8.841% | 4.504% | 8.273% |  | 10.5% | 1.852% | 7.047% | 5.053% | 5.346%
| PDM3 |  | 3.108% | -2.613% | -0.8394% | 6.655% | 0.5342% | 2.526% | -4.108% | -2.109% |  | 13.54% | -15.37% | -47.21% | -3.065% | 18.13%
| Z-Score (Absolute) |  | 0.79 | 0.36 | 0 | 1.5 | 0.28 | 0.67 | 0.65 | 0.26 | NA | 3.89 | 2.9 | 9.27 | 0.45 | 3.79
| Analytical results for MgO in OREAS 464 (Certified Value 1.4 wt.%) | [16 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.4157 | 1.443 | 1.443 | 1.3421 | N.A. | 1.4758 | 1.291 | 1.27 | 1.459 | 1.426 | 1.409 |  |  |  |  | 1
| 1 | 2 | 1.4037 | 1.459 | 1.459 | 1.3196 | N.A. | 1.4675 | 1.282 | 1.33 | 1.426 | 1.409 | 1.36 |  |  |  |  | 2
| 1 | 3 | 1.4277 | 1.476 | 1.459 | 1.3381 | N.A. | 1.4758 | 1.28 | 1.33 | 1.443 | 1.492 | 1.409 |  |  |  |  | 16
| 1 | 4 | 1.3969 | 1.459 | 1.459 | 1.3414 | N.A. | 1.4758 | 1.302 | 1.3 | 1.426 | 1.426 | 1.393 |  |  |  |  | 1.4021
| 1 | 5 | 1.4003 | 1.443 | 1.459 | 1.3359 | N.A. | 1.4841 | 1.285 | 1.33 | 1.409 | 1.459 | 1.409 |  |  |  |  | 39
| 1 | 6 | 1.3831 | 1.443 | 1.459 | 1.3374 | N.A. | 1.4841 | 1.298 | 1.34 | 1.393 | 1.459 | 1.409
| Mean |  | 1.4046 | 1.4538 | 1.4563 | 1.3358 |  | 1.4772 | 1.2897 | 1.3167 | 1.426 | 1.4452 | 1.3982
| Median |  | 1.402 | 1.451 | 1.459 | 1.3377 |  | 1.4758 | 1.288 | 1.33 | 1.426 | 1.4425 | 1.409
| Std Dev. |  | 0.015473 | 0.013393 | 0.006532 | 0.0082643 |  | 0.006248 | 0.0089144 | 0.026583 | 0.023478 | 0.030394 | 0.019763
| Rel.Std.Dev. |  | 1.102% | 0.9212% | 0.4485% | 0.6187% |  | 0.423% | 0.6912% | 2.019% | 1.646% | 2.103% | 1.413%
| PDM3 |  | 0.1769% | 3.69% | 3.869% | -4.731% |  | 5.356% | -8.018% | -6.093% | 1.705% | 3.072% | -0.28%
| Z-Score (Absolute) |  | 0.2 | 0.65 | 0.7 | 1.39 | NA | 1.06 | 2.19 | 1.56 | 0.17 | 0.51 | 0.17
| Analytical results for MnO in OREAS 464 (Certified Value 0.365 wt.%) | [16 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.39 | 0.362 | 0.37583 | N.A. | 0.4003 | 0.3347 | 0.23 | 0.4029 | 0.3667 | 0.3305 | [5 empty] | 1
| 1 | 2 | 0.39 | 0.362 | 0.37248 | N.A. | 0.3938 | 0.3297 | 0.35 | 0.3899 | 0.3615 | 0.3215 | [5 empty] | 3
| 1 | 3 | 0.39 | 0.362 | 0.35936 | N.A. | 0.4003 | 0.3318 | 0.34 | 0.3977 | 0.3719 | 0.3163 | [5 empty] | 16
| 1 | 4 | 0.39 | 0.362 | 0.35879 | N.A. | 0.4003 | 0.3348 | 0.26 | 0.3848 | 0.3667 | 0.3267 | [5 empty] | 0.36468
| 1 | 5 | 0.39 | 0.362 | 0.37194 | N.A. | 0.4003 | 0.3348 | 0.35 | 0.3796 | 0.3744 | 0.3267 | [5 empty] | 40
| 1 | 6 | 0.39 | 0.362 | 0.37184 | N.A. | 0.4003 | 0.3353 | 0.35 | 0.3822 | 0.3744 | 0.3331
| Mean |  | 0.39 | 0.362 | 0.36837 |  | 0.39922 | 0.33352 | 0.31333 | 0.38952 | 0.36927 | 0.3258
| Median |  | 0.39 | 0.362 | 0.37189 |  | 0.4003 | 0.33475 | 0.345 | 0.38735 | 0.3693 | 0.3267
| Std Dev. |  | 6.0809e-17 | 6.0809e-17 | 0.0073522 |  | 0.0026536 | 0.0022534 | 0.053914 | 0.009169 | 0.0051601 | 0.0060949
| Rel.Std.Dev. |  | 1.559e-14% | 1.68e-14% | 1.996% |  | 0.6647% | 0.6756% | 17.21% | 2.354% | 1.397% | 1.871%
| PDM3 |  | 6.944% | -0.734% | 1.014% |  | 9.471% | -8.545% | -14.08% | 6.811% | 1.259% | -10.66%
| Z-Score (Absolute) |  | 0.61 | 0.22 | 0 | NA | 1 | 1.13 | 0.67 | 0.68 | 0.03 | 1.37
| Analytical results for Mo in OREAS 464 (Certified Value 77 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
1
| 1 | 1 | 78 | 37.156 | N.A. | 80 | 72.7564 | 70 | 75 | 84 | [7 empty] | 1
| 1 | 2 | 76 | 37.2795 | N.A. | 80 | 74.9987 | 72 | 75 | 81 | [7 empty] | #N/A
| 1 | 3 | 77 | 36.5775 | N.A. | 80 | 71.7111 | 74 | 70 | 82 | [7 empty] | 16
| 1 | 4 | 78 | 37.8922 | N.A. | 85 | 72.88 | 76 | 75 | 84 | [7 empty] | 77.0732
| 1 | 5 | 77 | 36.3622 | N.A. | 80 | 75.1085 | 74 | 75 | 81 | [7 empty] | 41
| 1 | 6 | 76 | 37.841 | N.A. | 85 | 75.1795 | 70 | 80 | 82
| Mean |  | 77 | 37.1848 |  | 81.6667 | 73.7724 | 72.6667 | 75 | 82.3333
| Median |  | 77 | 37.2178 |  | 80 | 73.9393 | 73 | 75 | 82
| Std Dev. |  | 0.89443 | 0.63028 |  | 2.582 | 1.5064 | 2.4221 | 3.1623 | 1.3663
| Rel.Std.Dev. |  | 1.162% | 1.695% |  | 3.162% | 2.042% | 3.333% | 4.216% | 1.659%
| PDM3 |  | -0.09494% | -51.75% |  | 5.96% | -4.283% | -5.717% | -2.69% | 6.825%
| Z-Score (Absolute) |  | 0.58 | 10.93 | NA | 1.93 | 0.35 | 0.67 | 0 | 2.12
| Analytical results for Na in OREAS 464 (Indicative Value 0.135 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 11 | 14 | 18 | 20 | [11 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES | BF*OES | [11 empty] | 3
3
| 1 | 1 | 0.14 | N.A. | 0.12864 | 0.69 | [11 empty] | 1
| 1 | 2 | 0.14 | N.A. | 0.12856 | 0.67 | [11 empty] | 17
| 1 | 3 | 0.14 | N.A. | 0.1296 | 0.67 | [11 empty] | 16
| 1 | 4 | 0.15 | N.A. | 0.12908 | 0.66 | [11 empty] | 0.13501
| 1 | 5 | 0.14 | N.A. | 0.1293 | 0.63 | [11 empty] | 23
| 1 | 6 | 0.14 | N.A. | 0.12493 | 0.63
| Mean |  | 0.14167 |  | 0.12835 | 0.65833
| Median |  | 0.14 |  | 0.12886 | 0.665
| Std Dev. |  | 0.0040825 |  | 0.0017235 | 0.024014
| Rel.Std.Dev. |  | 2.882% |  | 1.343% | 3.648%
| PDM3 |  | 4.932% |  | -4.932% | 387.6%
| Z-Score (Absolute) |  | 0 | NA | 0.67 | 26.16
| Analytical results for Nb2O5 in OREAS 464 (Certified Value 2723 ppm) | [16 empty] | Certified
Nb2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 2746.6 | 2646 | 2646.4 | 2753.7 | 2596.4 | 2477.6 | 2574.9 | N.A. | 2203 | 1978 | 1406.2 | 2703.6 | 2932.5 |  |  | 1
| 1 | 2 | 2667.9 | 2661 | 2646.4 | 2782.3 | 2682.2 | 2746.6 | 2619.2 | N.A. | 2331.7 | 1961 | 2088.5 | 2718 | 2932.5 |  |  | 5
| 1 | 3 | 2861 | 2775 | 2617.8 | 2825.2 | 2560.6 | 2713.7 | 2597.8 | N.A. | 2260.2 | 1810 | 2131.4 | 2789.5 | 2889.6 |  |  | 16
| 1 | 4 | 2853.8 | 2604 | 2718 | 2818.1 | 2603.5 | 2798.1 | 2577.8 | N.A. | 2181.5 | 1988 | 1313.2 | 2732.3 | 2918.2 |  |  | 2722.8243
| 1 | 5 | 2760.9 | 2618 | 2675 | 2839.5 | 2517.7 | 2796.6 | 2589.2 | N.A. | 2338.9 | 2190 | 2117.1 | 2746.6 | 3275.8 |  |  | 42
| 1 | 6 | 2732.3 | 2675 | 2703.6 | 2753.7 | 2439 | 2778 | 2607.8 | N.A. | 2388.9 | 1785 | 2045.6 | 2818.1 | 2975.4
| Mean |  | 2770.4167 | 2663.1667 | 2667.8667 | 2795.4167 | 2566.5667 | 2718.4333 | 2594.45 |  | 2284.0333 | 1952 | 1850.3333 | 2751.35 | 2987.3333
| Median |  | 2753.75 | 2653.5 | 2660.7 | 2800.2 | 2578.5 | 2762.3 | 2593.5 |  | 2295.95 | 1969.5 | 2067.05 | 2739.45 | 2932.5
| Std Dev. |  | 74.5365 | 60.806 | 38.1298 | 37.41 | 82.8516 | 122.3286 | 17.2569 |  | 82.3735 | 146.1164 | 382.3041 | 44.0198 | 144.0139
| Rel.Std.Dev. |  | 2.69% | 2.283% | 1.429% | 1.338% | 3.228% | 4.5% | 0.6651% |  | 3.606% | 7.485% | 20.66% | 1.6% | 4.821%
| PDM3 |  | 1.748% | -2.191% | -2.018% | 2.666% | -5.739% | -0.1613% | -4.715% |  | -16.12% | -28.31% | -32.04% | 1.048% | 9.715%
| Z-Score (Absolute) |  | 0.69 | 0.02 | 0.02 | 0.85 | 0.65 | 0.66 | 0.47 | NA | 2.5 | 4.67 | 5.34 | 0.56 | 1.73
| Analytical results for Nd2O3 in OREAS 464 (Certified Value 1.16 wt.%) | [16 empty] | Certified
Nd2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.1571 | 1.1291 | > 1.2 | 1.1279 | 1.1442 | 1.1679 | 1.1366 | 1.1093 | 0.89264 | 1.18 | 1.1454 | 1.0422 | 1.0614 | 1.2131 | 1.2131 | 1
| 1 | 2 | 1.1477 | 1.1302 | > 1.2 | 1.1349 | 1.1664 | 1.178 | 1.1535 | 1.121 | 0.8941 | 1.18 | 1.1566 | 1.0572 | 1.1267 | 1.2364 | 1.2014 | #N/A
| 1 | 3 | 1.1781 | > 1.2 | 1.1594 | 1.1477 | 1.2014 | 1.2116 | 1.1637 | 1.088 | 0.89744 | 1.19 | 1.1389 | 1.0501 | 1.1617 | 1.1897 | 1.1781 | 16
| 1 | 4 | 1.1839 | > 1.2 | > 1.2 | 1.1431 | 1.1839 | 1.1765 | 1.1357 | 1.1142 | 0.89438 | 1.16 | 1.142 | 1.0467 | 1.1489 | 1.2014 | 1.2247 | 1.1593
| 1 | 5 | 1.1501 | 1.1641 | > 1.2 | 1.1501 | 1.1244 | 1.1951 | 1.1531 | 1.102 | 0.90066 | 1.18 | 1.1645 | 1.0675 | 1.1536 | 1.1897 | 1.2364 | 43
| 1 | 6 | 1.1582 | 1.0777 | > 1.2 | 1.1396 | 1.1956 | 1.1662 | 1.1656 | 1.0905 | 0.89088 | 1.16 | 1.1657 | 1.0643 | 1.1092 | 1.2131 | 1.2364
| Mean |  | 1.1625 | 1.1253 | 1.1594 | 1.1405 | 1.1693 | 1.1826 | 1.1514 | 1.1042 | 0.89502 | 1.175 | 1.1522 | 1.0547 | 1.1269 | 1.2072 | 1.215
| Median |  | 1.1577 | 1.1297 | 1.1594 | 1.1413 | 1.1751 | 1.1773 | 1.1533 | 1.1057 | 0.89424 | 1.18 | 1.151 | 1.0536 | 1.1378 | 1.2072 | 1.2189
| Std Dev. |  | 0.014978 | 0.035603 |  | 0.0082609 | 0.030297 | 0.017569 | 0.012837 | 0.01312 | 0.0035137 | 0.012247 | 0.011672 | 0.01001 | 0.037437 | 0.017688 | 0.022637
| Rel.Std.Dev. |  | 1.288% | 3.164% |  | 0.7243% | 2.591% | 1.486% | 1.115% | 1.188% | 0.3926% | 1.042% | 1.013% | 0.9491% | 3.322% | 1.465% | 1.863%
| PDM3 |  | 0.2734% | -2.938% | 0.004583% | -1.622% | 0.8599% | 2.002% | -0.6886% | -4.759% | -22.8% | 1.35% | -0.6176% | -9.03% | -2.796% | 4.13% | 4.8%
| Z-Score (Absolute) |  | 0.28 | 0.72 | 0.19 | 0.31 | 0.46 | 0.81 | 0.02 | 1.28 | 6.87 | 0.61 | 0 | 2.6 | 0.67 | 1.47 | 1.68
| Analytical results for NiO in OREAS 464 (Certified Value 104 ppm) | [16 empty] | Certified
NiO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 67.4 | N.A. | 130 | 113.589 | 64.9 | 102 | 127 | 96.7 | [7 empty] | 1
| 1 | 2 | 114.5 | N.A. | < 130 | 121.913 | 94.2 | 102 | 191 | 98 | [7 empty] | 6
| 1 | 3 | 63.6 | N.A. | < 130 | 108.8593 | 96.7 | 102 | 127 | 87.8 | [7 empty] | 16
| 1 | 4 | 52.2 | N.A. | < 130 | 114.06 | 75.1 | 102 | 127 | 90.3 | [7 empty] | 104.3283
| 1 | 5 | 78.9 | N.A. | < 130 | 113.8055 | 96.7 | 102 | 127 | 92.9 | [7 empty] | 44
| 1 | 6 | 62.4 | N.A. | < 130 | 116.0043 | 92.9 | 102 | 127 | 104.3
| Mean |  | 73.1667 |  | 130 | 114.7052 | 86.75 | 102 | 137.6667 | 95
| Median |  | 65.5 |  | 130 | 113.9327 | 93.55 | 102 | 127 | 94.8
| Std Dev. |  | 22.0073 |  |  | 4.2494 | 13.4497 | 0 | 26.1279 | 5.9451
| Rel.Std.Dev. |  | 30.08% |  |  | 3.705% | 15.5% | 0% | 18.98% | 6.258%
| PDM3 |  | -29.87% |  | 24.61% | 9.946% | -16.85% | -2.232% | 31.96% | -8.941%
| Z-Score (Absolute) |  | 1.5 | NA | 1.33 | 0.67 | 0.56 | 0 | 1.33 | 0.35
| Analytical results for P2O5 in OREAS 464 (Certified Value 3.23 wt.%) | [16 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 3.3114 | 3.208 | 3.025 | 3.3136 | N.A. | 3.2312 | 3.273 | 2.864 | 3.254 | 3.117 | [5 empty] | 1
| 1 | 2 | 3.2816 | 3.254 | 3.117 | 3.2736 | N.A. | 3.2655 | 3.295 | 2.91 | 3.346 | 3.139 | [5 empty] | 7
| 1 | 3 | 3.2999 | 3.185 | 3.139 | 3.1664 | N.A. | 3.4374 | 3.29 | 2.796 | 3.437 | 3.025 | [5 empty] | 16
| 1 | 4 | 3.222 | 3.208 | 3.071 | 3.1586 | N.A. | 3.1968 | 3.264 | 2.887 | 3.437 | 3.139 | [5 empty] | 3.2271
| 1 | 5 | 3.2953 | 3.208 | 3.048 | 3.2688 | N.A. | 3.2197 | 3.314 | 2.796 | 3.437 | 3.117 | [5 empty] | 45
| 1 | 6 | 3.3091 | 3.139 | 3.094 | 3.2512 | N.A. | 3.2999 | 3.295 | 2.842 | 3.346 | 3.071
| Mean |  | 3.2866 | 3.2003 | 3.0823 | 3.2387 |  | 3.2751 | 3.2885 | 2.8492 | 3.3762 | 3.1013
| Median |  | 3.2976 | 3.208 | 3.0825 | 3.26 |  | 3.2483 | 3.2925 | 2.853 | 3.3915 | 3.117
| Std Dev. |  | 0.033383 | 0.037559 | 0.042763 | 0.062503 |  | 0.087379 | 0.017762 | 0.047025 | 0.074628 | 0.044889
| Rel.Std.Dev. |  | 1.016% | 1.174% | 1.387% | 1.93% |  | 2.668% | 0.5401% | 1.65% | 2.21% | 1.447%
| PDM3 |  | 1.842% | -0.8292% | -4.486% | 0.3593% |  | 1.487% | 1.903% | -11.71% | 4.619% | -3.897%
| Z-Score (Absolute) |  | 0.65 | 0.52 | 2.12 | 0 | NA | 0.05 | 0.67 | 5.27 | 1.86 | 1.86
| Analytical results for PbO in OREAS 464 (Certified Value 187 ppm) | [16 empty] | Certified
PbO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 189.6 | 179.9 | 320 | 177.2337 | 95.9 | < 320 | 183 | 200.4 | [7 empty] | 1
| 1 | 2 | 186.4 | 209 | 320 | 177.4966 | 94.8 | < 320 | 172 | 229.4 | [7 empty] | 8
| 1 | 3 | 182 | 201.4 | 320 | 176.0576 | 94.8 | < 320 | 172 | 187.4 | [7 empty] | 16
| 1 | 4 | 181 | 218.7 | 220 | 176.1396 | 100.2 | < 320 | 172 | 189.6 | [7 empty] | 186.9196
| 1 | 5 | 188.5 | 217.6 | 320 | 177.9122 | 96.9 | < 320 | 172 | 202.5 | [7 empty] | 46
| 1 | 6 | 185.3 | 192.8 | 320 | 177.1806 | 88.3 | < 320 | 172 | 193.9
| Mean |  | 185.4667 | 203.2333 | 303.3333 | 177.0034 | 95.15 |  | 173.8333 | 200.5333
| Median |  | 185.85 | 205.2 | 320 | 177.2072 | 95.35 |  | 172 | 197.15
| Std Dev. |  | 3.4396 | 15.0701 | 40.8248 | 0.74756 | 3.9063 |  | 4.4907 | 15.3179
| Rel.Std.Dev. |  | 1.855% | 7.415% | 13.46% | 0.4223% | 4.105% |  | 2.583% | 7.639%
| PDM3 |  | -0.7773% | 8.728% | 62.28% | -5.305% | -49.1% |  | -7.001% | 7.283%
| Z-Score (Absolute) |  | 0 | 1.06 | 7.14 | 0.5 | 5.38 | NA | 0.67 | 0.55
| Analytical results for Pr6O11 in OREAS 464 (Certified Value 3138 ppm) | [16 empty] | Certified
Pr6O11 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 3165 | 2996 | 3117.2 | > 1200 | 3045 | 3216.4526 | 3072.94 | 3299.47 | 2485.9596 | 3140 | 3089.97 | 3398.9041 | 2863.434 | 3141.3 | 2923.8 | 1
| 1 | 2 | 3141 | 3008 | 3117.2 | > 1200 | 3069 | 3223.9355 | 3209.82 | 3287.75 | 2372.9024 | 3120 | 3127.43 | 3398.6805 | 3032.582 | 3177.6 | 2875.5 | #N/A
| 1 | 3 | 3214 | 3153 | 2972.2 | > 1200 | 3105 | 3301.8654 | 3156.54 | 3308.05 | 2580.2879 | 3150 | 3076.68 | 3349.3321 | 3153.402 | 3093 | 2827.2 | 16
| 1 | 4 | 3238 | 3008 | 3238 | > 1200 | 3057 | 3188.4418 | 3149.17 | 3320.25 | 2476.2904 | 3070 | 3077.29 | 3358.7888 | 3129.238 | 3153.4 | 2911.8 | 3137.8453
| 1 | 5 | 3141 | 2996 | 3117.2 | > 1200 | 2972 | 3281.346 | 3157.27 | 3294.76 | 2425.1906 | 3120 | 3138.9 | 3418.4118 | 3092.992 | 3153.4 | 2984.3 | 47
| 1 | 6 | 3178 | 3020 | 3141.3 | > 1200 | 3141 | 3191.196 | 3170.8 | 3268.42 | 2412.6554 | 3100 | 3172.13 | 3405.4027 | 3008.418 | 3105.1 | 2935.9
| Mean |  | 3179.5 | 3030.1667 | 3117.1833 |  | 3064.8333 | 3233.8729 | 3152.7567 | 3296.45 | 2458.8811 | 3116.6667 | 3113.7333 | 3388.2533 | 3046.6777 | 3137.3 | 2909.75
| Median |  | 3171.5 | 3008 | 3117.2 |  | 3063 | 3220.194 | 3156.905 | 3297.115 | 2450.7405 | 3120 | 3108.7 | 3398.7923 | 3062.787 | 3147.35 | 2917.8
| Std Dev. |  | 39.4702 | 60.8421 | 85.0864 |  | 57.454 | 47.2529 | 44.7214 | 17.7446 | 72.7188 | 28.7518 | 38.7218 | 27.6001 | 105.4439 | 32.118 | 53.7228
| Rel.Std.Dev. |  | 1.241% | 2.008% | 2.73% |  | 1.875% | 1.461% | 1.418% | 0.5383% | 2.957% | 0.9225% | 1.244% | 0.8146% | 3.461% | 1.024% | 1.846%
| PDM3 |  | 1.327% | -3.432% | -0.6585% |  | -2.327% | 3.06% | 0.4752% | 5.055% | -21.64% | -0.6749% | -0.7684% | 7.98% | -2.905% | -0.01738% | -7.269%
| Z-Score (Absolute) |  | 0.5 | 1.15 | 0.18 | NA | 0.59 | 1.01 | 0.24 | 1.6 | 6.32 | 0.1 | 0.13 | 2.47 | 0.76 | 0.1 | 2.06
| Analytical results for Rb in OREAS 464 (Certified Value 4.39 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  | 2
2
| 1 | 1 | 3.5 | 4 | 4.2 | 4.2 | 4 | 4.6974 | 3.5 | 4.7 | 6.9576 | N.A. | 4 | 3.3 | 5.5 | 5.2 |  | 1
| 1 | 2 | 4.5 | 4.3 | 4.4 | 4.3 | 4 | 4.7745 | 3.8 | 5.3 | 6.0712 | N.A. | 4 | 3.4 | 5.5 | 5.2 |  | #N/A
| 1 | 3 | 4.8 | 4.1 | 4.4 | 4.1 | 4 | 4.2735 | 4.5 | 4.7 | 5.8513 | N.A. | 4 | 3.4 | 5.5 | 3.9 |  | 16
| 1 | 4 | 5.2 | 4.6 | 4.4 | 4.2 | 4 | 4.184 | 3.8 | 4.7 | 5.9953 | N.A. | 5 | 3.3 | 5.5 | 4.6 |  | 4.3867
| 1 | 5 | 6.3 | 4.6 | 4.4 | 4 | 4 | 4.7381 | 4.4 | 5.5 | 6.1281 | N.A. | 5 | 3.4 | 5.5 | 4.4 |  | 48
| 1 | 6 | 4.5 | 3.8 | 4.4 | 3.9 | 5 | 4.3331 | 3.8 | 5.1 | 5.9305 | N.A. | 4 | 3.2 | 5.5 | 4.8
| Mean |  | 4.8 | 4.2333 | 4.3667 | 4.1167 | 4.1667 | 4.5001 | 3.9667 | 5 | 6.1557 |  | 4.3333 | 3.3333 | 5.5 | 4.6833
| Median |  | 4.65 | 4.2 | 4.4 | 4.15 | 4 | 4.5152 | 3.8 | 4.9 | 6.0333 |  | 4 | 3.35 | 5.5 | 4.7
| Std Dev. |  | 0.9252 | 0.3266 | 0.08165 | 0.1472 | 0.40825 | 0.26456 | 0.39328 | 0.35214 | 0.40498 |  | 0.5164 | 0.08165 | 0 | 0.49967
| Rel.Std.Dev. |  | 19.28% | 7.715% | 1.87% | 3.576% | 9.798% | 5.879% | 9.915% | 7.043% | 6.579% |  | 11.92% | 2.449% | 0% | 10.67%
| PDM3 |  | 9.422% | -3.496% | -0.4561% | -6.155% | -5.015% | 2.586% | -9.575% | 13.98% | 40.33% |  | -1.216% | -24.01% | 25.38% | 6.763%
| Z-Score (Absolute) |  | 0.36 | 0.36 | 0 | 0.67 | 0.54 | 0.36 | 1.08 | 1.71 | 4.39 | NA | 0.09 | 2.79 | 3.06 | 0.85
| Analytical results for Re in OREAS 464 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 17 | 22 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | [12 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 1 | [12 empty] | 18
| 1 | 3 | 0.2 | < 0.1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | [12 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | [12 empty] | 24
| 1 | 6 | < 0.1 | < 0.1 | < 1
| Mean |  | 0.2
| Median |  | 0.2
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 464 (Certified Value 1028 ppm) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 0
0
| 1 | 1 | 980 | 1000 | 1000 | 1000 | 1000 | [10 empty] | 1
| 1 | 2 | 940 | 1000 | 1100 | 900 | 1200 | [10 empty] | #N/A
| 1 | 3 | 950 | 1000 | 1000 | 900 | 1200 | [10 empty] | 16
| 1 | 4 | 980 | 1100 | 1000 | 1100 | 1000 | [10 empty] | 1028.3333
| 1 | 5 | 970 | 1100 | 1100 | 1000 | 1200 | [10 empty] | 49
| 1 | 6 | 930 | 1000 | 1100 | 900 | 1200
| Mean |  | 958.3333 | 1033.3333 | 1050 | 966.6667 | 1133.3333
| Median |  | 960 | 1000 | 1050 | 950 | 1200
| Std Dev. |  | 21.3698 | 51.6398 | 54.7723 | 81.6497 | 103.2796
| Rel.Std.Dev. |  | 2.23% | 4.997% | 5.216% | 8.447% | 9.113%
| PDM3 |  | -6.807% | 0.4862% | 2.107% | -5.997% | 10.21%
| Z-Score (Absolute) |  | 0.76 | 0 | 0.17 | 0.67 | 1.01
| Analytical results for Sb in OREAS 464 (Certified Value 1.8 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.8 | 2.9 | 2 | 1.623 | 2 | 1.4 | [9 empty] | 1
| 1 | 2 | 2 | 2.7 | 2 | 1.6574 | 2 | 1.5 | [9 empty] | #N/A
| 1 | 3 | 1.8 | 2.8 | 2 | 1.4668 | 2 | 1.5 | [9 empty] | 16
| 1 | 4 | 2 | 2.8 | 2 | 1.3877 | 2 | 1.5 | [9 empty] | 1.8015
| 1 | 5 | 1.9 | 3 | 2 | 1.4881 | 2 | 1.7 | [9 empty] | 50
| 1 | 6 | 2.2 | 2.7 | 2 | 1.4204 | 2 | 1.7
| Mean |  | 1.95 | 2.8167 | 2 | 1.5073 | 2 | 1.55
| Median |  | 1.95 | 2.8 | 2 | 1.4775 | 2 | 1.5
| Std Dev. |  | 0.15166 | 0.1169 | 0 | 0.10935 | 0 | 0.12247
| Rel.Std.Dev. |  | 7.777% | 4.15% | 0% | 7.255% | 0% | 7.902%
| PDM3 |  | 8.246% | 56.36% | 11.02% | -16.33% | 11.02% | -13.96%
| Z-Score (Absolute) |  | 0.07 | 2.52 | 0.07 | 1.4 | 0.07 | 1.27
| Analytical results for Sc in OREAS 464 (Indicative Value 152 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | [11 empty] | 0
0
| 1 | 1 | 150 | 150 | 140 | 170 | [11 empty] | 1
| 1 | 2 | 152 | 150 | 140 | 160 | [11 empty] | 19
| 1 | 3 | 150 | 150 | 140 | 160 | [11 empty] | 16
| 1 | 4 | 151 | 150 | 140 | 170 | [11 empty] | 151.875
| 1 | 5 | 151 | 150 | 140 | 170 | [11 empty] | 25
| 1 | 6 | 151 | 150 | 140 | 170
| Mean |  | 150.8333 | 150 | 140 | 166.6667
| Median |  | 151 | 150 | 140 | 170
| Std Dev. |  | 0.75277 | 0 | 0 | 5.164
| Rel.Std.Dev. |  | 0.4991% | 0% | 0% | 3.098%
| PDM3 |  | -0.6859% | -1.235% | -7.819% | 9.739%
| Z-Score (Absolute) |  | 0.05 | 0.05 | 1.3 | 2.02
| Analytical results for Se in OREAS 464 (Indicative Value 40 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 1
1
| 1 | 1 | < 20 | 40 | [13 empty] | 1
| 1 | 2 | < 20 | 40 | [13 empty] | 15
| 1 | 3 | < 20 | 40 | [13 empty] | 16
| 1 | 4 | < 20 | 40 | [13 empty] | 40
| 1 | 5 | < 20 | 40 | [13 empty] | 21
| 1 | 6 | < 20 | 40
| Mean |  |  | 40
| Median |  |  | 40
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for SiO2 in OREAS 464 (Certified Value 20.36 wt.%) | [16 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 20.876 | 19.595 | 20.7456 | N.A. | 20.6 | 19.4954 | 26.9 | 19.766 | [7 empty] | 1
| 1 | 2 | 20.8464 | 20.194 | 20.7489 | N.A. | 20.451 | 19.6452 | 28.9 | 19.574 | [7 empty] | 9
| 1 | 3 | 22.0183 | 19.916 | 20.6923 | N.A. | 20.322 | 19.6131 | 28.4 | 19.959 | [7 empty] | 16
| 1 | 4 | 21.3414 | 20.002 | 20.94 | N.A. | 20.536 | 19.795 | 27.5 | 19.766 | [7 empty] | 20.3638
| 1 | 5 | 21.5318 | 20.258 | 21.0163 | N.A. | 20.558 | 19.7736 | 27.8 | 19.959 | [7 empty] | 51
| 1 | 6 | 21.3202 | 20.087 | 20.7463 | N.A. | 20.729 | 19.7201 | 29.1 | 19.959
| Mean |  | 21.3223 | 20.0087 | 20.8149 |  | 20.5327 | 19.6737 | 28.1 | 19.8305
| Median |  | 21.3308 | 20.0445 | 20.7476 |  | 20.547 | 19.6826 | 28.1 | 19.8625
| Std Dev. |  | 0.4369 | 0.23766 | 0.13049 |  | 0.13768 | 0.11236 | 0.85088 | 0.15726
| Rel.Std.Dev. |  | 2.049% | 1.188% | 0.6269% |  | 0.6706% | 0.5711% | 3.028% | 0.793%
| PDM3 |  | 4.707% | -1.744% | 2.215% |  | 0.8293% | -3.389% | 37.99% | -2.619%
| Z-Score (Absolute) |  | 0.76 | 0.5 | 0.27 | NA | 0 | 0.82 | 7.26 | 0.67
| Analytical results for Sm2O3 in OREAS 464 (Certified Value 1738 ppm) | [16 empty] | Certified
Sm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1727.8 | 1716.2 | 1814.77 | > 1200 | 1733.6 | 1765.444 | 1709.13 | 1766.77 | 1506.5848 | 1700 | 1734.18 | 1771.3763 | 1623.44 | 1681.4 | 1727.8 | 1
| 1 | 2 | 1698.8 | 1756.8 | 1797.38 | > 1200 | 1756.8 | 1760.1404 | 1763.87 | 1797.38 | 1425.0051 | 1700 | 1747.52 | 1740.3488 | 1727.8 | 1739.4 | 1704.6 | #N/A
| 1 | 3 | 1733.6 | 1780 | 1716.21 | > 1200 | 1820.6 | 1804.512 | 1724.44 | 1740.44 | 1399.3564 | 1700 | 1723.17 | 1722.0019 | 1785.78 | 1693 | 1658.2 | 16
| 1 | 4 | 1756.8 | 1681.4 | 1861.16 | > 1200 | 1780 | 1725.0972 | 1718.41 | 1789.84 | 1523.1211 | 1700 | 1718.53 | 1728.8907 | 1611.84 | 1681.4 | 1727.8 | 1737.5013
| 1 | 5 | 1693 | 1710.4 | 1797.38 | > 1200 | 1704.6 | 1772.2039 | 1725.72 | 1789.03 | 1490.928 | 1700 | 1763.17 | 1745.5417 | 1751 | 1681.4 | 1751 | 52
| 1 | 6 | 1704.6 | 1710.4 | 1791.58 | > 1200 | 1797.4 | 1777.4802 | 1750.88 | 1780.33 | 1488.7619 | 1700 | 1775.35 | 1711.6116 | 1693.02 | 1739.4 | 1762.6
| Mean |  | 1719.1 | 1725.8667 | 1796.4133 |  | 1765.5 | 1767.4796 | 1732.075 | 1777.2983 | 1472.2929 | 1700 | 1743.6533 | 1736.6285 | 1698.8133 | 1702.6667 | 1722
| Median |  | 1716.2 | 1713.3 | 1797.38 |  | 1768.4 | 1768.824 | 1725.08 | 1784.68 | 1489.8449 | 1700 | 1740.85 | 1734.6197 | 1710.41 | 1687.2 | 1727.8
| Std Dev. |  | 24.5389 | 35.8788 | 46.8762 |  | 42.5816 | 25.8766 | 20.8635 | 20.8651 | 48.8545 | 0 | 22.5434 | 20.9778 | 69.8655 | 28.806 | 37.2286
| Rel.Std.Dev. |  | 1.427% | 2.079% | 2.609% |  | 2.412% | 1.464% | 1.205% | 1.174% | 3.318% | 0% | 1.293% | 1.208% | 4.113% | 1.692% | 2.162%
| PDM3 |  | -1.059% | -0.6696% | 3.391% |  | 1.611% | 1.725% | -0.3123% | 2.29% | -15.26% | -2.158% | 0.3541% | -0.05024% | -2.227% | -2.005% | -0.8922%
| Z-Score (Absolute) |  | 0.24 | 0.08 | 1.65 | NA | 0.89 | 0.94 | 0.08 | 1.18 | 6.26 | 0.71 | 0.36 | 0.19 | 0.74 | 0.64 | 0.17
| Analytical results for SnO2 in OREAS 464 (Certified Value 58 ppm) | [16 empty] | Certified
SnO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS |  | 0
1
| 1 | 1 | 60.9 | 57.1 | 60.9 | 57.1 | 58.4 | 58.4 | 45.7 | N.A. | 53.3 | 32.2651 | 63.5 | < 63 | 63 | 43.2 |  | 1
| 1 | 2 | 58.4 | 60.9 | 62.2 | 57.1 | 60.9 | 59.7 | 45.7 | N.A. | 54.6 | 28.4472 | 52.1 | < 63 | 63 | 41.9 |  | 10
| 1 | 3 | 62.2 | 58.4 | 59.7 | 58.4 | 58.4 | 58.4 | 47 | N.A. | 50.8 | 29.6071 | 54.6 | < 63 | 63 | 41.9 |  | 16
| 1 | 4 | 62.2 | 55.9 | 60.9 | 57.1 | 58.4 | 58.4 | 47 | N.A. | 54.6 | 30.9735 | 39.4 | < 63 | 51 | 44.4 |  | 58.2323
| 1 | 5 | 59.7 | 57.1 | 60.9 | 58.4 | 55.9 | 58.4 | 45.7 | N.A. | 54.6 | 33.3768 | 54.6 | < 63 | 63 | 45.7 |  | 53
| 1 | 6 | 59.7 | 57.1 | 62.2 | 58.4 | 53.3 | 58.4 | 48.2 | N.A. | 54.6 | 32.7607 | 53.3 | < 63 | 63 | 44.4
| Mean |  | 60.5167 | 57.75 | 61.1333 | 57.75 | 57.55 | 58.6167 | 46.55 |  | 53.75 | 31.2384 | 52.9167 |  | 61 | 43.5833
| Median |  | 60.3 | 57.1 | 60.9 | 57.75 | 58.4 | 58.4 | 46.35 |  | 54.6 | 31.6193 | 53.95 |  | 63 | 43.8
| Std Dev. |  | 1.525 | 1.7341 | 0.94798 | 0.71204 | 2.6144 | 0.53072 | 1.0291 |  | 1.5359 | 1.9218 | 7.7613 |  | 4.899 | 1.525
| Rel.Std.Dev. |  | 2.52% | 3.003% | 1.551% | 1.233% | 4.543% | 0.9054% | 2.211% |  | 2.857% | 6.152% | 14.67% |  | 8.031% | 3.499%
| PDM3 |  | 3.923% | -0.8283% | 4.982% | -0.8283% | -1.172% | 0.66% | -20.06% |  | -7.697% | -46.36% | -9.128% |  | 4.753% | -25.16%
| Z-Score (Absolute) |  | 0.61 | 0.02 | 0.74 | 0.02 | 0.02 | 0.2 | 2.36 | NA | 0.83 | 5.61 | 0.44 | NA | 0.79 | 2.99
| Analytical results for SrO in OREAS 464 (Certified Value 2437 ppm) | [16 empty] | Certified
SrO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 2542.6 | 2359.3 | 2347.5 | 2388.9 | 2400.7 | 2427.9 | 2362.24 | N.A. | 2480 | 3158.1084 | 1998.59 | 2282.4 | 2519 |  |  | 1
| 1 | 2 | 2495.3 | 2377 | 2459.8 | 2412.5 | 2471.6 | 2452.7 | 2348.17 | N.A. | 2480 | 3195.3727 | 2022.25 | 2400.7 | 2543 |  |  | 1
| 1 | 3 | 2649 | 2448 | 2347.5 | 2448 | 2359.3 | 2412.5 | 2330.2 | N.A. | 2370 | 3096.9628 | 2069.55 | 2412.5 | 2507 |  |  | 16
| 1 | 4 | 2625.4 | 2530.8 | 2448 | 2436.2 | 2388.9 | 2445.6 | 2348.88 | N.A. | 2480 | 3142.2221 | 1750.25 | 2388.9 | 2543 |  |  | 2436.6409
| 1 | 5 | 2566.2 | 2459.8 | 2507.1 | 2436.2 | 2412.5 | 2453.9 | 2349.23 | N.A. | 2480 | 3210.1158 | 2010.42 | 2258.8 | 2554 |  |  | 54
| 1 | 6 | 2518.9 | 2258.8 | 2471.6 | 2400.7 | 2341.5 | 2442.1 | 2381.05 | N.A. | 2480 | 3241.6922 | 1963.12 | 2412.5 | 2531
| Mean |  | 2566.2333 | 2405.6167 | 2430.25 | 2420.4167 | 2395.75 | 2439.1167 | 2353.295 |  | 2461.6667 | 3174.079 | 1969.03 | 2359.3 | 2532.8333
| Median |  | 2554.4 | 2412.5 | 2453.9 | 2424.35 | 2394.8 | 2443.85 | 2349.055 |  | 2480 | 3176.7406 | 2004.505 | 2394.8 | 2537
| Std Dev. |  | 60.3018 | 94.7864 | 67.0807 | 23.2542 | 45.5563 | 16.0515 | 17.006 |  | 44.9073 | 52.0393 | 112.6261 | 69.6627 | 17.4174
| Rel.Std.Dev. |  | 2.35% | 3.94% | 2.76% | 0.9608% | 1.902% | 0.6581% | 0.7226% |  | 1.824% | 1.64% | 5.72% | 2.953% | 0.6877%
| PDM3 |  | 5.318% | -1.273% | -0.2623% | -0.6658% | -1.678% | 0.1016% | -3.421% |  | 1.027% | 30.26% | -19.19% | -3.174% | 3.948%
| Z-Score (Absolute) |  | 1.82 | 0.25 | 0.06 | 0.06 | 0.38 | 0.18 | 0.93 | NA | 0.5 | 9.67 | 5.33 | 0.85 | 1.39
| Analytical results for Ta2O5 in OREAS 464 (Certified Value 29.7 ppm) | [16 empty] | Certified
Ta2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 1
1
| 1 | 1 | 30.53 | 31.14 | 31.99 | 32.6 | 31.75 | 21.6757 | 24.18 | 592.2 | N.A. | 20.8 | 18.1661 | 23.81 | 23.81 | 30.89 |  | 1
| 1 | 2 | 29.79 | 30.77 | 32.73 | 32.97 | 31.5 | 22.0768 | 30.65 | 628.9 | N.A. | 23.2 | 12.8691 | 35.53 | 25.64 | 29.79 |  | 2
| 1 | 3 | 30.16 | 31.02 | 31.38 | 32.73 | 32.97 | 20.8218 | 31.14 | 525.1 | N.A. | 22 | 13.3381 | 35.66 | 23.81 | 35.66 |  | 16
| 1 | 4 | 32.48 | 33.7 | 34.31 | 33.46 | 31.63 | 21.9769 | 31.14 | 650.8 | N.A. | 20.8 | 13.7562 | 26.74 | 26.86 | 30.41 |  | 29.695
| 1 | 5 | 30.77 | 31.75 | 32.36 | 33.21 | 30.89 | 27.5007 | 31.26 | 639.9 | N.A. | 23.2 | 19.1627 | 35.66 | 26.25 | 31.75 |  | 55
| 1 | 6 | 30.53 | 30.89 | 33.34 | 32.48 | 32.97 | 31.9149 | 31.02 | 660.6 | N.A. | 23.2 | 17.8875 | 34.92 | 26.86 | 30.89
| Mean |  | 30.71 | 31.545 | 32.685 | 32.9083 | 31.9517 | 24.3278 | 29.8983 | 616.25 |  | 22.2 | 15.8633 | 32.0533 | 25.5383 | 31.565
| Median |  | 30.53 | 31.08 | 32.545 | 32.85 | 31.69 | 22.0268 | 31.08 | 634.4 |  | 22.6 | 15.8219 | 35.225 | 25.945 | 30.89
| Std Dev. |  | 0.93276 | 1.1095 | 1.0357 | 0.37669 | 0.84253 | 4.417 | 2.8093 | 50.5306 |  | 1.1798 | 2.8308 | 5.3386 | 1.4131 | 2.1071
| Rel.Std.Dev. |  | 3.037% | 3.517% | 3.169% | 1.145% | 2.637% | 18.16% | 9.396% | 8.2% |  | 5.315% | 17.85% | 16.66% | 5.533% | 6.675%
| PDM3 |  | 3.418% | 6.23% | 10.07% | 10.82% | 7.6% | -18.07% | 0.6849% | 1975% |  | -25.24% | -46.58% | 7.942% | -14% | 6.298%
| Z-Score (Absolute) |  | 0.41 | 0 | 0.56 | 0.67 | 0.2 | 4.32 | 0.25 | 289.15 | NA | 4.63 | 7.75 | 1.07 | 2.97 | 0.01
| Analytical results for Tb4O7 in OREAS 464 (Certified Value 63 ppm) | [16 empty] | Certified
Tb4O7 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 58.575 | 62.34 | 66.691 | 58.81 | 59.75 | 65.107 | 61.63 | 62.57 | 52.9824 | 70 | 63.51 | 77.7844 | 57.399 | 64.22 | 60.221 | 1
| 1 | 2 | 58.457 | 61.28 | 64.691 | 59.751 | 61.63 | 64.7109 | 62.93 | 64.93 | 51.8522 | 70 | 64.81 | 78.2639 | 61.045 | 68.45 | 60.104 | 3
| 1 | 3 | 59.163 | 62.34 | 63.75 | 60.574 | 62.22 | 66.65 | 62.1 | 61.87 | 53.4267 | 70 | 63.63 | 76.205 | 62.809 | 65.4 | 55.517 | 16
| 1 | 4 | 61.398 | 63.99 | 66.455 | 59.751 | 61.87 | 65.3782 | 62.57 | 64.93 | 53.8163 | 70 | 64.57 | 76.1192 | 62.339 | 66.57 | 61.162 | 63.2658
| 1 | 5 | 60.104 | 61.52 | 65.75 | 59.633 | 59.4 | 68.4005 | 62.22 | 64.81 | 52.6424 | 70 | 65.51 | 77.6267 | 62.221 | 66.1 | 60.81 | 56
| 1 | 6 | 58.34 | 62.46 | 64.338 | 59.28 | 62.1 | 66.8193 | 62.81 | 63.4 | 53.1945 | 70 | 65.28 | 76.3681 | 59.516 | 66.1 | 61.868
| Mean |  | 59.3395 | 62.3217 | 65.2792 | 59.6332 | 61.1617 | 66.1776 | 62.3767 | 63.7517 | 52.9857 | 70 | 64.5517 | 77.0612 | 60.8882 | 66.14 | 59.947
| Median |  | 58.869 | 62.34 | 65.2205 | 59.692 | 61.75 | 66.0141 | 62.395 | 64.105 | 53.0885 | 70 | 64.69 | 76.9974 | 61.633 | 66.1 | 60.5155
| Std Dev. |  | 1.2013 | 0.95193 | 1.1973 | 0.58571 | 1.2504 | 1.3793 | 0.48784 | 1.3385 | 0.68273 | 0 | 0.83072 | 0.93702 | 2.081 | 1.397 | 2.2643
| Rel.Std.Dev. |  | 2.024% | 1.527% | 1.834% | 0.9822% | 2.044% | 2.084% | 0.7821% | 2.1% | 1.289% | 0% | 1.287% | 1.216% | 3.418% | 2.112% | 3.777%
| PDM3 |  | -6.206% | -1.492% | 3.182% | -5.742% | -3.326% | 4.603% | -1.405% | 0.768% | -16.25% | 10.64% | 2.033% | 21.81% | -3.758% | 4.543% | -5.246%
| Z-Score (Absolute) |  | 0.75 | 0.01 | 0.71 | 0.67 | 0.3 | 0.93 | 0 | 0.34 | 2.31 | 1.87 | 0.53 | 3.61 | 0.37 | 0.92 | 0.38
| Analytical results for Te in OREAS 464 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 1 | < 1 | [13 empty] | 1
| 1 | 2 | < 1 | < 1 | [13 empty] | 16
| 1 | 3 | < 1 | < 1 | [13 empty] | 16
| 1 | 4 | < 1 | < 1 | [13 empty] | < 1
| 1 | 5 | < 1 | < 1 | [13 empty] | 22
| 1 | 6 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for ThO2 in OREAS 464 (Certified Value 600 ppm) | [16 empty] | Certified
ThO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 595.12 | 583.7 | 617.88 | 590.57 | 591.71 | 620.0341 | 626.3 | 600.36 | 417.2655 | 611.1 | 597.97 | 423.4492 | 474.5 | 572.4 | 561 | 1
| 1 | 2 | 590.57 | 592.8 | 614.47 | 604.22 | 608.78 | 615.6042 | 626.87 | 596.49 | 411.9612 | 625.8 | 591.71 | 429.4594 | 538.23 | 665.7 | 547.3 | #N/A
| 1 | 3 | 612.19 | 605.4 | 591.71 | 613.33 | 573.5 | 626.619 | 618.9 | 595.58 | 417.4615 | 624.7 | 586.59 | 423.5498 | 546.19 | 615.6 | 538.2 | 16
| 1 | 4 | 636.09 | 603.1 | 640.64 | 607.64 | 588.29 | 605.324 | 613.33 | 610.03 | 417.4509 | 608.8 | 588.86 | 428.3857 | 566.67 | 645.2 | 561 | 600.1434
| 1 | 5 | 615.6 | 595.1 | 607.64 | 605.36 | 631.53 | 626.5365 | 626.41 | 598.31 | 419.3105 | 617.9 | 598.54 | 434.2316 | 534.81 | 633.8 | 562.1 | 57
| 1 | 6 | 600.81 | 576.9 | 616.74 | 607.64 | 592.85 | 610.2297 | 632.79 | 593.53 | 418.0892 | 612.2 | 598.54 | 432.4555 | 521.16 | 653.2 | 567.8
| Mean |  | 608.3967 | 592.8333 | 614.8467 | 604.7933 | 597.7767 | 617.3913 | 624.1 | 599.05 | 416.9231 | 616.75 | 593.7017 | 428.5885 | 530.26 | 630.9833 | 556.2333
| Median |  | 606.5 | 593.95 | 615.605 | 606.5 | 592.28 | 617.8192 | 626.355 | 597.4 | 417.4562 | 615.05 | 594.84 | 428.9225 | 536.52 | 639.5 | 561
| Std Dev. |  | 16.6422 | 11.0068 | 15.9089 | 7.6424 | 20.004 | 8.6703 | 6.876 | 5.8633 | 2.5443 | 7.2423 | 5.3483 | 4.4587 | 31.163 | 33.3992 | 11.1234
| Rel.Std.Dev. |  | 2.735% | 1.857% | 2.587% | 1.264% | 3.346% | 1.404% | 1.102% | 0.9788% | 0.6102% | 1.174% | 0.9008% | 1.04% | 5.877% | 5.293% | 2%
| PDM3 |  | 1.375% | -1.218% | 2.45% | 0.7748% | -0.3944% | 2.874% | 3.992% | -0.1822% | -30.53% | 2.767% | -1.073% | -28.59% | -11.64% | 5.139% | -7.317%
| Z-Score (Absolute) |  | 0.36 | 0.24 | 0.6 | 0.22 | 0.05 | 0.7 | 0.95 | 0 | 6.94 | 0.67 | 0.2 | 6.5 | 2.2 | 1.22 | 1.49
| Analytical results for TiO2 in OREAS 464 (Certified Value 3.26 wt.%) | [16 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 3.3078 | 3.236 | 3.086 | 3.2012 | N.A. | 3.353 | 3.1622 | 1.918 | 3.5214 | 3.269 | 3.32 |  |  |  |  | 1
| 1 | 2 | 3.3012 | 3.269 | 3.136 | 3.0747 | N.A. | 3.32 | 3.169 | 3.053 | 3.4296 | 3.203 | 3.286 |  |  |  |  | 4
| 1 | 3 | 3.3445 | 3.253 | 3.153 | 3.1337 | N.A. | 3.303 | 3.1565 | 3.019 | 3.4813 | 3.286 | 3.303 |  |  |  |  | 16
| 1 | 4 | 3.2761 | 3.253 | 3.153 | 3.0789 | N.A. | 3.336 | 3.1987 | 2.185 | 3.3846 | 3.269 | 3.303 |  |  |  |  | 3.2593
| 1 | 5 | 3.3512 | 3.253 | 3.153 | 3.1105 | N.A. | 3.336 | 3.1776 | 3.069 | 3.3379 | 3.336 | 3.286 |  |  |  |  | 58
| 1 | 6 | 3.3395 | 3.219 | 3.153 | 3.163 | N.A. | 3.37 | 3.1728 | 3.086 | 3.3596 | 3.303 | 3.269
| Mean |  | 3.3201 | 3.2472 | 3.139 | 3.127 |  | 3.3363 | 3.1728 | 2.7217 | 3.4191 | 3.2777 | 3.2945
| Median |  | 3.3237 | 3.253 | 3.153 | 3.1221 |  | 3.336 | 3.1709 | 3.036 | 3.4071 | 3.2775 | 3.2945
| Std Dev. |  | 0.029609 | 0.017302 | 0.02684 | 0.04934 |  | 0.023619 | 0.014748 | 0.5264 | 0.071789 | 0.044361 | 0.01783
| Rel.Std.Dev. |  | 0.8918% | 0.5328% | 0.8551% | 1.578% |  | 0.7079% | 0.4648% | 19.34% | 2.1% | 1.353% | 0.5412%
| PDM3 |  | 1.865% | -0.3715% | -3.69% | -4.059% |  | 2.364% | -2.653% | -16.49% | 4.903% | 0.5643% | 1.081%
| Z-Score (Absolute) |  | 0.47 | 0.13 | 1.02 | 1.12 | NA | 0.61 | 0.74 | 3.14 | 1.29 | 0.13 | 0.26
| Analytical results for Tl in OREAS 464 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 14 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 17
| 1 | 3 | < 0.5 | N.A. | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | < 0.5
| 1 | 5 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 23
| 1 | 6 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5
| Mean |  |  |  | 0.1
| Median |  |  |  | 0.1
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Tm2O3 in OREAS 464 (Certified Value 4.06 ppm) | [16 empty] | Certified
Tm2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 3.883 | 3.997 | 4.077 | 4.34 | 3.94 | 3.8224 | 4.11 | 4.91 | 2.7018 | < 50 | 4 | 3.3301 | 3.963 | 3.88 | 3.803 | 1
| 1 | 2 | 3.86 | 4.1 | 4.237 | 4.363 | 3.872 | 3.6602 | 4.11 | 4.45 | 2.6289 | < 50 | 4.11 | 3.3781 | 4.26 | 4.11 | 4.077 | 5
| 1 | 3 | 3.872 | 3.986 | 4.249 | 4.351 | 4.169 | 4.0965 | 4.11 | 4.57 | 2.5546 | < 50 | 4 | 3.2982 | 4.351 | 4.11 | 3.746 | 16
| 1 | 4 | 4.009 | 4.066 | 4.454 | 4.409 | 3.86 | 3.7999 | 4 | 4.23 | 2.746 | < 50 | 3.88 | 3.3069 | 4.192 | 4.11 | 3.906 | 4.0622
| 1 | 5 | 3.769 | 4.089 | 4.351 | 4.386 | 3.906 | 4.1249 | 4.11 | 4.8 | 2.6524 | < 50 | 4 | 3.3147 | 4.249 | 3.88 | 4.192 | 59
| 1 | 6 | 3.872 | 3.895 | 4.637 | 4.363 | 4.112 | 3.803 | 4.11 | 4.34 | 2.6803 | < 50 | 3.88 | 3.2871 | 4.157 | 4.11 | 3.952
| Mean |  | 3.8775 | 4.0222 | 4.3342 | 4.3687 | 3.9765 | 3.8845 | 4.0917 | 4.55 | 2.6607 |  | 3.9783 | 3.3192 | 4.1953 | 4.0333 | 3.946
| Median |  | 3.872 | 4.0315 | 4.3 | 4.363 | 3.923 | 3.8127 | 4.11 | 4.51 | 2.6663 |  | 4 | 3.3108 | 4.2205 | 4.11 | 3.929
| Std Dev. |  | 0.076761 | 0.078178 | 0.19452 | 0.025001 | 0.13131 | 0.18474 | 0.044907 | 0.2642 | 0.065861 |  | 0.087274 | 0.032345 | 0.13172 | 0.11877 | 0.16719
| Rel.Std.Dev. |  | 1.98% | 1.944% | 4.488% | 0.5723% | 3.302% | 4.756% | 1.098% | 5.807% | 2.475% |  | 2.194% | 0.9745% | 3.14% | 2.945% | 4.237%
| PDM3 |  | -4.547% | -0.9861% | 6.694% | 7.544% | -2.11% | -4.376% | 0.7248% | 12.01% | -34.5% |  | -2.065% | -18.29% | 3.277% | -0.7112% | -2.861%
| Z-Score (Absolute) |  | 0.76 | 0.11 | 1.7 | 1.88 | 0.12 | 0.59 | 0.47 | 2.8 | 6.82 | NA | 0.11 | 3.47 | 0.99 | 0.18 | 0.28
| Analytical results for U3O8 in OREAS 464 (Certified Value 20.7 ppm) | [16 empty] | Certified
U3O8 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 19.93 | 20.16 | 21.23 | 20.695 | 20.28 | 21.1152 | 20.64 | 20.99 | 16.6317 | 20 | 20.64 | 20.4454 | 11.709 | 18.28 | 21.108 | 1
| 1 | 2 | 20.05 | 20.4 | 20.87 | 20.577 | 20.4 | 21.5791 | 20.87 | 21.34 | 16.2107 | 20 | 20.64 | 20.742 | 13.561 | 18.87 | 21.108 | 6
| 1 | 3 | 20.05 | 20.99 | 20.4 | 20.872 | 21.34 | 21.1728 | 20.28 | 20.99 | 17.2108 | 20 | 20.64 | 20.426 | 13.797 | 18.28 | 19.693 | 16
| 1 | 4 | 20.28 | 21.11 | 21.46 | 20.872 | 20.4 | 20.7915 | 20.28 | 20.87 | 17.465 | 20 | 21.23 | 20.7002 | 10.896 | 18.28 | 20.99 | 20.7145
| 1 | 5 | 19.57 | 20.4 | 20.64 | 21.344 | 19.81 | 21.7706 | 20.4 | 21.34 | 16.7548 | 20 | 21.23 | 20.8196 | 13.325 | 18.28 | 21.697 | 60
| 1 | 6 | 20.28 | 20.28 | 20.52 | 20.518 | 21.46 | 21.3831 | 20.87 | 20.87 | 17.4506 | 20 | 21.23 | 20.7564 | 12.971 | 17.69 | 20.872
| Mean |  | 20.0267 | 20.5567 | 20.8533 | 20.813 | 20.615 | 21.3021 | 20.5567 | 21.0667 | 16.9539 | 20 | 20.935 | 20.6483 | 12.7098 | 18.28 | 20.9113
| Median |  | 20.05 | 20.4 | 20.755 | 20.7835 | 20.4 | 21.278 | 20.52 | 20.99 | 16.9828 | 20 | 20.935 | 20.7211 | 13.148 | 18.28 | 21.049
| Std Dev. |  | 0.26357 | 0.39419 | 0.41769 | 0.29852 | 0.64689 | 0.35075 | 0.27602 | 0.21842 | 0.50394 | 0 | 0.32316 | 0.16916 | 1.1528 | 0.37315 | 0.66132
| Rel.Std.Dev. |  | 1.316% | 1.918% | 2.003% | 1.434% | 3.138% | 1.647% | 1.343% | 1.037% | 2.972% | 0% | 1.544% | 0.8192% | 9.07% | 2.041% | 3.162%
| PDM3 |  | -3.32% | -0.7618% | 0.6703% | 0.4756% | -0.4802% | 2.837% | -0.7618% | 1.7% | -18.15% | -3.449% | 1.065% | -0.3196% | -38.64% | -11.75% | 0.9503%
| Z-Score (Absolute) |  | 0.88 | 0.09 | 0.35 | 0.29 | 0 | 1.02 | 0.09 | 0.67 | 5.46 | 0.85 | 0.47 | 0.05 | 11.79 | 3.49 | 0.8
| Analytical results for V2O5 in OREAS 464 (Certified Value 412 ppm) | [16 empty] | Certified
V2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 396.3 | 424.9 | 440.9 | 414.2 | 374.9 | 421.3 | 399.9 | N.A. | 360 | 430.023 | 469.5 | 446 | 383.8 |  |  | 1
| 1 | 2 | 383.8 | 421.3 | 432 | 412.4 | 389.2 | 428.4 | 401.7 | N.A. | 360 | 427.8249 | 469.5 | 357 | 374.9 |  |  | 7
| 1 | 3 | 407 | 421.3 | 437.4 | 423.1 | 374.9 | 439.2 | 392.7 | N.A. | 360 | 435.3958 | 471.3 | 446 | 364.2 |  |  | 16
| 1 | 4 | 410.6 | 462.4 | 444.5 | 423.1 | 382 | 424.9 | 385.6 | N.A. | 360 | 425.9646 | 490.9 | 446 | 382 |  |  | 412.2966
| 1 | 5 | 394.5 | 426.7 | 465.9 | 423.1 | 439.2 | 423.1 | 369.5 | N.A. | 360 | 432.2727 | 474.9 | 446 | 396.3 |  |  | 61
| 1 | 6 | 387.4 | 396.3 | 442.7 | 408.8 | 423.1 | 412.4 | 392.7 | N.A. | 360 | 440.9671 | 482 | 357 | 387.4
| Mean |  | 396.6 | 425.4833 | 443.9 | 417.45 | 397.2167 | 424.8833 | 390.35 |  | 360 | 432.0747 | 476.35 | 416.3333 | 381.4333
| Median |  | 395.4 | 423.1 | 441.8 | 418.65 | 385.6 | 424 | 392.7 |  | 360 | 431.1478 | 473.1 | 446 | 382.9
| Std Dev. |  | 10.5559 | 21.2195 | 11.6504 | 6.4289 | 27.2926 | 8.8226 | 11.7333 |  | 0 | 5.4726 | 8.5449 | 45.9594 | 10.9746
| Rel.Std.Dev. |  | 2.662% | 4.987% | 2.625% | 1.54% | 6.871% | 2.076% | 3.006% |  | 0% | 1.267% | 1.794% | 11.04% | 2.877%
| PDM3 |  | -3.807% | 3.198% | 7.665% | 1.25% | -3.658% | 3.053% | -5.323% |  | -12.68% | 4.797% | 15.54% | 0.9791% | -7.486%
| Z-Score (Absolute) |  | 0.68 | 0.04 | 0.76 | 0.01 | 0.66 | 0.26 | 0.89 | NA | 2.01 | 0.51 | 1.99 | 0.01 | 1.19
| Analytical results for WO3 in OREAS 464 (Certified Value < 7 ppm) | [16 empty] | Certified
WO3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 17 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 5 | 3.8 | 5 | 5 | 5 | 3.8 | 6.3 | < 13 | 6.3 | < 6.3 | 5 |  |  |  |  | 1
| 1 | 2 | 5 | 3.8 | 6.3 | 3.8 | 6.3 | 5 | 3.8 | < 13 | 7.6 | < 6.3 | 6.3 |  |  |  |  | 8
| 1 | 3 | 6.3 | 5 | 5 | 3.8 | 5 | 3.8 | 3.8 | < 13 | 7.6 | < 6.3 | 6.3 |  |  |  |  | 16
| 1 | 4 | 5 | 5 | 5 | 3.8 | 6.3 | 3.8 | 3.8 | 13 | 6.3 | < 6.3 | 6.3 |  |  |  |  | < 7
| 1 | 5 | 5 | 5 | 5 | 5 | 3.8 | 3.8 | 3.8 | 25 | 7.6 | < 6.3 | 6.3 |  |  |  |  | 62
| 1 | 6 | 3.8 | 5 | 6.3 | 3.8 | 3.8 | 3.8 | 3.8 | < 13 | 7.6 | < 6.3 | 6.3
| Mean |  | 5.0167 | 4.6 | 5.4333 | 4.2 | 5.0333 | 4 | 4.2167 | 19 | 7.1667 |  | 6.0833
| Median |  | 5 | 5 | 5 | 3.8 | 5 | 3.8 | 3.8 | 19 | 7.6 |  | 6.3
| Std Dev. |  | 0.79099 | 0.61968 | 0.67132 | 0.61968 | 1.1183 | 0.4899 | 1.0206 | 8.4853 | 0.67132 |  | 0.53072
| Rel.Std.Dev. |  | 15.77% | 13.47% | 12.36% | 14.75% | 22.22% | 12.25% | 24.2% | 44.66% | 9.367% |  | 8.724%
PDM3
| Z-Score (Absolute) |  | 0 | 0.27 | 0.27 | 0.54 | 0 | 0.81 | 0.81 | 8.9 | 1.35 | NA | 0.81
| Analytical results for Y2O3 in OREAS 464 (Certified Value 570 ppm) | [16 empty] | Certified
Y2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 560 | 575.3 | 585.4 | 581.61 | 568.9 | 626.2815 | 570.19 | 555.96 | 550 | 539.71 | 471 | 612.09 | 594.3 | 518.12 |  | 1
| 1 | 2 | 548.6 | 591.8 | 585.4 | 585.42 | 575.3 | 605.2858 | 578.19 | 549.87 | 540 | 548.6 | 484 | 726.38 | 627.3 | 510.5 |  | #N/A
| 1 | 3 | 563.8 | 588 | 580.3 | 587.96 | 589.2 | 594.2582 | 567.26 | 549.1 | 560 | 537.8 | 490 | 739.08 | 615.9 | 502.88 |  | 16
| 1 | 4 | 575.3 | 560 | 605.7 | 590.5 | 576.5 | 581.8712 | 573.99 | 547.33 | 540 | 532.09 | 481 | 601.93 | 600.7 | 515.58 |  | 569.6169
| 1 | 5 | 552.4 | 571.5 | 596.9 | 593.04 | 571.5 | 603.799 | 577.68 | 597.74 | 550 | 553.04 | 494 | 736.54 | 594.3 | 528.28 |  | 63
| 1 | 6 | 557.5 | 567.6 | 591.8 | 582.88 | 604.5 | 601.8374 | 573.87 | 561.42 | 550 | 556.85 | 489 | 721.3 | 612.1 | 527.01
| Mean |  | 559.6 | 575.7 | 590.9167 | 586.9017 | 580.9833 | 602.2222 | 573.53 | 560.2367 | 548.3333 | 544.6817 | 484.8333 | 689.5533 | 607.4333 | 517.0617
| Median |  | 558.75 | 573.4 | 588.6 | 586.69 | 575.9 | 602.8182 | 573.93 | 552.915 | 550 | 544.155 | 486.5 | 723.84 | 606.4 | 516.85
| Std Dev. |  | 9.3972 | 12.1688 | 9.2597 | 4.4299 | 13.4825 | 14.6182 | 4.2364 | 19.1011 | 7.5277 | 9.6323 | 8.1833 | 64.3477 | 13.2548 | 9.7192
| Rel.Std.Dev. |  | 1.679% | 2.114% | 1.567% | 0.7548% | 2.321% | 2.427% | 0.7386% | 3.409% | 1.373% | 1.768% | 1.688% | 9.332% | 2.182% | 1.88%
| PDM3 |  | -1.759% | 1.068% | 3.739% | 3.034% | 1.995% | 5.724% | 0.687% | -1.647% | -3.736% | -4.378% | -14.88% | 21.06% | 6.639% | -9.226%
| Z-Score (Absolute) |  | 0.42 | 0.03 | 0.46 | 0.34 | 0.05 | 0.77 | 0.03 | 0.61 | 0.74 | 0.84 | 2.51 | 3.22 | 0.92 | 1.61
| Analytical results for Yb2O3 in OREAS 464 (Certified Value 17.9 ppm) | [16 empty] | Certified
Yb2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 17.88 | 17.251 | 16.74 | 17.65 | 18.219 | 16.0842 | 17.88 | 18.67 | 17.9616 | < 100 | 19.81 | 18.4197 | 17.08 | 15.94 | 22.66 | 1
| 1 | 2 | 17.76 | 17.878 | 16.4 | 17.422 | 18.561 | 18.0287 | 17.88 | 18.56 | 16.6085 | < 100 | 19.24 | 18.8355 | 20.38 | 17.65 | 21.86 | 9
| 1 | 3 | 17.76 | 18.276 | 16.63 | 17.878 | 17.536 | 17.7281 | 18.11 | 18.11 | 16.7967 | < 100 | 19.24 | 18.3635 | 20.5 | 18.22 | 21.75 | 16
| 1 | 4 | 18.22 | 19.927 | 16.85 | 17.08 | 18.105 | 16.8138 | 18.67 | 18.45 | 17.2893 | < 100 | 19.59 | 18.37 | 17.31 | 17.08 | 22.43 | 17.919
| 1 | 5 | 17.08 | 18.048 | 16.28 | 17.935 | 17.024 | 17.9095 | 18.11 | 21.75 | 17.1996 | < 100 | 19.93 | 18.4232 | 21.07 | 17.08 | 22.89 | 64
| 1 | 6 | 17.31 | 17.08 | 17.54 | 17.764 | 15.999 | 15.9855 | 18.11 | 18.56 | 17.2251 | < 100 | 20.15 | 18.1669 | 20.27 | 15.94 | 22.66
| Mean |  | 17.6683 | 18.0767 | 16.74 | 17.6215 | 17.574 | 17.0916 | 18.1267 | 19.0167 | 17.1801 |  | 19.66 | 18.4298 | 19.435 | 16.985 | 22.375
| Median |  | 17.76 | 17.963 | 16.685 | 17.707 | 17.8205 | 17.2709 | 18.11 | 18.56 | 17.2123 |  | 19.7 | 18.3949 | 20.325 | 17.08 | 22.545
| Std Dev. |  | 0.41 | 1.0177 | 0.44529 | 0.32194 | 0.94411 | 0.92387 | 0.28904 | 1.3529 | 0.46879 |  | 0.37234 | 0.21993 | 1.7584 | 0.91319 | 0.46617
| Rel.Std.Dev. |  | 2.321% | 5.63% | 2.66% | 1.827% | 5.372% | 5.405% | 1.595% | 7.114% | 2.729% |  | 1.894% | 1.193% | 9.047% | 5.376% | 2.083%
| PDM3 |  | -1.399% | 0.8802% | -6.579% | -1.66% | -1.925% | -4.617% | 1.159% | 6.126% | -4.123% |  | 9.716% | 2.851% | 8.461% | -5.212% | 24.87%
| Z-Score (Absolute) |  | 0.18 | 0.25 | 1.15 | 0.23 | 0.28 | 0.78 | 0.18 | 0.66 | 0.69 | NA | 1.9 | 0.62 | 1.67 | 0.9 | 4.74
| Analytical results for ZnO in OREAS 464 (Certified Value 1533 ppm) | [16 empty] | Certified
ZnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 14 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | N.A. | 1490 | 1682.44 | 1593 | 1432 | 1406.6 | [9 empty] | 1
| 1 | 2 | N.A. | 1490 | 1700.166 | 1556 | 1432 | 1394.2 | [9 empty] | 10
| 1 | 3 | N.A. | 1490 | 1745.1402 | 1581 | 1494 | 1381.7 | [9 empty] | 16
| 1 | 4 | N.A. | 1490 | 1773.1051 | 1593 | 1494 | 1381.7 | [9 empty] | 1533.1462
| 1 | 5 | N.A. | 1490 | 1729.4873 | 1531 | 1556 | 1381.7 | [9 empty] | 65
| 1 | 6 | N.A. | 1490 | 1750.3193 | 1506 | 1556 | 1381.7
| Mean |  |  | 1490 | 1730.1096 | 1560 | 1494 | 1387.9333
| Median |  |  | 1490 | 1737.3138 | 1568.5 | 1494 | 1381.7
| Std Dev. |  |  | 0 | 33.6225 | 35.7547 | 55.4545 | 10.4224
| Rel.Std.Dev. |  |  | 0% | 1.943% | 2.292% | 3.712% | 0.7509%
| PDM3 |  |  | -2.814% | 12.85% | 1.752% | -2.553% | -9.472%
| Z-Score (Absolute) |  | NA | 0 | 2.4 | 0.67 | 0 | 1.07
| Analytical results for ZrO2 in OREAS 464 (Certified Value 624 ppm) | [16 empty] | Certified
ZrO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 621 | 621 | 634.9 | 629.5 | 621 | 670 | 597.1 | N.A. | 602.5 | 602 | 1485.88 |  |  |  |  | 1
| 1 | 2 | 594 | 635 | 648.4 | 624.1 | 608 | 656.5 | 595.7 | N.A. | 601.1 | 480 | 2647.57 |  |  |  |  | 11
| 1 | 3 | 635 | 621 | 657.8 | 630.8 | 608 | 641.6 | 574.1 | N.A. | 588.9 | 446 | 2661.08 |  |  |  |  | 16
| 1 | 4 | 648 | 621 | 656.5 | 645.7 | 608 | 644.3 | 586.2 | N.A. | 601.1 | 468 | 1458.86 |  |  |  |  | 623.7038
| 1 | 5 | 608 | 621 | 665.9 | 634.9 | 608 | 644.3 | 593 | N.A. | 598.4 | 500 | 2688.09 |  |  |  |  | 66
| 1 | 6 | 621 | 635 | 672.7 | 630.8 | 635 | 629.5 | 603.8 | N.A. | 606.5 | 406 | 2566.52
| Mean |  | 621.1667 | 625.6667 | 656.0333 | 632.6333 | 614.6667 | 647.7 | 591.65 |  | 599.75 | 483.6667 | 2251.3333
| Median |  | 621 | 621 | 657.15 | 630.8 | 608 | 644.3 | 594.35 |  | 601.1 | 474 | 2607.045
| Std Dev. |  | 19.0936 | 7.2296 | 13.2918 | 7.2833 | 11.2368 | 13.9018 | 10.3272 |  | 5.9383 | 66.3013 | 604.8021
| Rel.Std.Dev. |  | 3.074% | 1.155% | 2.026% | 1.151% | 1.828% | 2.146% | 1.745% |  | 0.9901% | 13.71% | 26.86%
| PDM3 |  | -0.4068% | 0.3147% | 5.183% | 1.432% | -1.449% | 3.847% | -5.139% |  | -3.841% | -22.45% | 261%
| Z-Score (Absolute) |  | 0.06 | 0.06 | 0.91 | 0.25 | 0.25 | 0.68 | 0.9 | NA | 0.67 | 3.94 | 45.78
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 464 (Certified Value 4.43 ppm) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 5.11 | 5.07 | 2.82 | 3.98 | 3.64 | 3.9122 | 4.3 | 4.8 | < 0.01 | 4.95 | 4 | 4.5 |  |  | 1
| 1 | 2 | 4.5 | 4.91 | 1.72 | 3.84 | 3.2 | 4.2153 | 4.3 | 4.5 | < 0.01 | 4.62 | 4 | 4.5 |  |  | 15
| 1 | 3 | 4.25 | 4.27 | 2.55 | 3.96 | 3.26 | 4.1407 | 4.7 | 4.5 | < 0.01 | 5.03 | 4 | 4.5 |  |  | 16
| 1 | 4 | 4.67 | 4.43 | 1.94 | 4.09 | 3.3 | 4.2482 | 4.4 | 4.9 | < 0.01 | 4.88 | 6 | 4.5 |  |  | 4.431
| 1 | 5 | 4.66 | 4.67 | 3.03 | 4.19 | 3.35 | 3.9543 | 4.2 | 4.5 | < 0.01 | 4.96 | 4 | 4.5 |  |  | 68
| 1 | 6 | 4.84 | 4.85 | 3.39 | 3.91 | 3.28 | 4.1732 | 4.1 | 5 | < 0.01 | 4.79 | 4 | 4.5
| Mean |  | 4.6717 | 4.7 | 2.575 | 3.995 | 3.3383 | 4.1073 | 4.3333 | 4.7 |  | 4.8717 | 4.3333 | 4.5
| Median |  | 4.665 | 4.76 | 2.685 | 3.97 | 3.29 | 4.1569 | 4.3 | 4.65 |  | 4.915 | 4 | 4.5
| Std Dev. |  | 0.29253 | 0.30404 | 0.64283 | 0.12629 | 0.15575 | 0.14031 | 0.20656 | 0.22804 |  | 0.1477 | 0.8165 | 0
| Rel.Std.Dev. |  | 6.262% | 6.469% | 24.96% | 3.161% | 4.665% | 3.416% | 4.767% | 4.852% |  | 3.032% | 18.84% | 0%
| PDM3 |  | 5.431% | 6.071% | -41.89% | -9.84% | -24.66% | -7.305% | -2.204% | 6.071% |  | 9.945% | -2.204% | 1.557%
| Z-Score (Absolute) |  | 0.67 | 0.73 | 3.5 | 0.67 | 2.1 | 0.45 | 0 | 0.73 | NA | 1.07 | 0.66 | 0.33
| Analytical results for Al in OREAS 464 (Certified Value 4.47 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 4.31 | 4.47 | 4.55 | 4.82 | 4.38 | 4.4999 | 4.45 | 4.58 | 4.45 | 4.5104 | 4.1042 | 4.61 | 4.576 | 4.7 | 1
| 1 | 2 | 4.33 | 4.42 | 4.49 | 4.81 | 4.61 | 4.4001 | 4.51 | 4.55 | 4.39 | 4.6046 | 4.1117 | 4.79 | 4.583 | 4.65 | #N/A
| 1 | 3 | 4.25 | 4.42 | 4.49 | 4.77 | 4.27 | 4.4034 | 4.54 | 4.55 | 4.48 | 4.4891 | 4.1659 | 4.65 | 4.528 | 4.83 | 16
| 1 | 4 | 4.28 | 4.37 | 4.49 | 4.97 | 4.61 | 4.5123 | 4.47 | 4.58 | 4.38 | 4.3918 | 4.1773 | 4.89 | 4.553 | 4.78 | 4.47
| 1 | 5 | 4.31 | 4.47 | 4.53 | 4.91 | 4.5 | 4.4466 | 4.51 | 4.47 | 4.42 | 4.3676 | 4.1645 | 4.7 | 4.586 | 4.7 | 69
| 1 | 6 | 4.37 | 4.49 | 4.56 | 4.79 | 4.38 | 4.6427 | 4.48 | 4.61 | 4.45 | 4.3029 | 4.1643 | 4.82 | 4.582 | 4.81
| Mean |  | 4.3083 | 4.44 | 4.5183 | 4.845 | 4.4583 | 4.4841 | 4.4933 | 4.5567 | 4.4283 | 4.4444 | 4.148 | 4.7433 | 4.568 | 4.745
| Median |  | 4.31 | 4.445 | 4.51 | 4.815 | 4.44 | 4.4732 | 4.495 | 4.565 | 4.435 | 4.4405 | 4.1644 | 4.745 | 4.579 | 4.74
| Std Dev. |  | 0.041191 | 0.044721 | 0.032506 | 0.07791 | 0.13819 | 0.090754 | 0.03266 | 0.048028 | 0.038687 | 0.11019 | 0.031476 | 0.10764 | 0.022952 | 0.071764
| Rel.Std.Dev. |  | 0.9561% | 1.007% | 0.7194% | 1.608% | 3.1% | 2.024% | 0.7269% | 1.054% | 0.8736% | 2.479% | 0.7588% | 2.269% | 0.5025% | 1.512%
| PDM3 |  | -3.616% | -0.6709% | 1.082% | 8.39% | -0.2607% | 0.3168% | 0.5223% | 1.939% | -0.9319% | -0.5724% | -7.204% | 6.115% | 2.193% | 6.152%
| Z-Score (Absolute) |  | 1.9 | 0.51 | 0.31 | 3.74 | 0.32 | 0.05 | 0.05 | 0.71 | 0.63 | 0.47 | 3.58 | 2.68 | 0.83 | 2.69
| Analytical results for As in OREAS 464 (Certified Value 32.9 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 36.2 | 35.8 | 28.1 | 34.4 | 52.6 | 32.0107 | 28 | 37 | 32.1 | 84.5 | < 50 | 70 |  |  | 1
| 1 | 2 | 35.6 | 35 | 29.1 | 32.3 | 51.8 | 30.2213 | 29 | 39 | 30.9 | 77.6 | < 50 | 74 |  |  | 16
| 1 | 3 | 36.3 | 31.6 | 27.5 | 32.4 | 52.5 | 32.2249 | 27 | 46 | 31.9 | 80.2 | < 50 | 72 |  |  | 16
| 1 | 4 | 36 | 28.5 | 28.9 | 31.9 | 54.9 | 30.5103 | 29 | 43 | 32 | 80.1 | < 50 | 69 |  |  | 32.9362
| 1 | 5 | 35.4 | 34 | 27 | 34.1 | 54.8 | 29.6059 | 28 | 44 | 30.9 | 82.7 | < 50 | 73 |  |  | 70
| 1 | 6 | 36.7 | 32.5 | 26.8 | 33.8 | 51.4 | 32.2638 | 29 | 46 | 31.4 | 80 | < 50 | 72
| Mean |  | 36.0333 | 32.9 | 27.9 | 33.15 | 53 | 31.1395 | 28.3333 | 42.5 | 31.5333 | 80.85 |  | 71.6667
| Median |  | 36.1 | 33.25 | 27.8 | 33.1 | 52.55 | 31.2605 | 28.5 | 43.5 | 31.65 | 80.15 |  | 72
| Std Dev. |  | 0.4761 | 2.6548 | 0.9654 | 1.071 | 1.5007 | 1.1655 | 0.8165 | 3.7283 | 0.5465 | 2.4089 |  | 1.8619
| Rel.Std.Dev. |  | 1.321% | 8.069% | 3.46% | 3.231% | 2.831% | 3.743% | 2.882% | 8.772% | 1.733% | 2.98% |  | 2.598%
| PDM3 |  | 9.403% | -0.1099% | -15.29% | 0.6492% | 60.92% | -5.455% | -13.98% | 29.04% | -4.259% | 145.5% |  | 117.6%
| Z-Score (Absolute) |  | 0.4 | 0.03 | 0.73 | 0 | 2.78 | 0.28 | 0.67 | 1.31 | 0.23 | 6.68 | NA | 5.39
| Analytical results for Ba in OREAS 464 (Certified Value 1648 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 1650 | 1730 | 1700 | 1770 | 1630 | 1672.0617 | 1509 | 1663 | 1650 | 1580.4 | 1636.93 | 1495 | 1655 | 1570 | 1
| 1 | 2 | 1640 | 1680 | 1690 | 1780 | 1720 | 1730.6561 | 1525 | 1662 | 1600 | 1529.1 | 1623.25 | 1583 | 1665 | 1590 | #N/A
| 1 | 3 | 1650 | 1700 | 1680 | 1760 | 1590 | 1708.8388 | 1511 | 1667 | 1650 | 1700.8 | 1641.51 | 1553 | 1655 | 1590 | 16
| 1 | 4 | 1660 | 1680 | 1720 | 1840 | 1700 | 1696.2428 | 1525 | 1646 | 1600 | 1565.4 | 1653.42 | 1649 | 1655 | 1610 | 1648.1311
| 1 | 5 | 1640 | 1400 | 1710 | 1820 | 1660 | 1697.2616 | 1497 | 1638 | 1670 | 2797.8 | 1634.77 | 1566 | 1660 | 1640 | 71
| 1 | 6 | 1670 | 1730 | 1700 | 1780 | 1640 | 1665.6526 | 1484 | 1670 | 1680 | 1558.8 | 1651.02 | 1618 | 1655 | 1630
| Mean |  | 1651.6667 | 1653.3333 | 1700 | 1791.6667 | 1656.6667 | 1695.1189 | 1508.5 | 1657.6667 | 1641.6667 | 1788.7167 | 1640.15 | 1577.3333 | 1657.5 | 1605
| Median |  | 1650 | 1690 | 1700 | 1780 | 1650 | 1696.7522 | 1510 | 1662.5 | 1650 | 1572.9 | 1639.22 | 1574.5 | 1655 | 1600
| Std Dev. |  | 11.6905 | 126.1216 | 14.1421 | 31.2517 | 47.6095 | 23.9035 | 16.0219 | 12.7227 | 34.3026 | 497.897 | 11.1447 | 53.5188 | 4.1833 | 26.6458
| Rel.Std.Dev. |  | 0.7078% | 7.628% | 0.8319% | 1.744% | 2.874% | 1.41% | 1.062% | 0.7675% | 2.089% | 27.84% | 0.6795% | 3.393% | 0.2524% | 1.66%
| PDM3 |  | 0.2145% | 0.3156% | 3.147% | 8.709% | 0.5179% | 2.851% | -8.472% | 0.5786% | -0.3922% | 8.53% | -0.4843% | -4.296% | 0.5685% | -2.617%
| Z-Score (Absolute) |  | 0.04 | 0.77 | 0.71 | 2.14 | 0.04 | 0.64 | 2.26 | 0.05 | 0.19 | 1.05 | 0.22 | 1.19 | 0.05 | 0.76
| Analytical results for Be in OREAS 464 (Certified Value 15.2 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 13.05 | 15.35 | 13.6 | 17.15 | 11.45 | 15.3338 | 15.2 | 16.5 | 15.1 | 15.06 | 13.5936 | 19 | 15.6 |  | 1
| 1 | 2 | 13.25 | 15.6 | 13.6 | 16.6 | 10.75 | 15.38 | 15.7 | 16.5 | 15.2 | 15.47 | 14.1298 | 19 | 16 |  | #N/A
| 1 | 3 | 13.4 | 13.85 | 14.7 | 17.15 | 10.85 | 15.3764 | 15.2 | 17.6 | 15.4 | 15.84 | 14.1889 | 19 | 16 |  | 16
| 1 | 4 | 13.8 | 14.3 | 13.9 | 17.8 | 10.6 | 15.12 | 15.3 | 17.3 | 14.9 | 15.91 | 13.9878 | 19.5 | 15.7 |  | 15.2434
| 1 | 5 | 13.6 | 14.45 | 14.05 | 17.25 | 10.75 | 15.0475 | 14.9 | 17.1 | 15.3 | 16.68 | 14.0684 | 19 | 16.4 |  | 72
| 1 | 6 | 13.95 | 15.35 | 13.6 | 16.95 | 10.2 | 15.1839 | 14.9 | 17 | 15.5 | 15.32 | 14.0353 | 19.5 | 16.2
| Mean |  | 13.5083 | 14.8167 | 13.9083 | 17.15 | 10.7667 | 15.2403 | 15.2 | 17 | 15.2333 | 15.7133 | 14.0006 | 19.1667 | 15.9833
| Median |  | 13.5 | 14.9 | 13.75 | 17.15 | 10.75 | 15.2588 | 15.2 | 17.05 | 15.25 | 15.655 | 14.0518 | 19 | 16
| Std Dev. |  | 0.33973 | 0.70969 | 0.43176 | 0.3937 | 0.40579 | 0.14254 | 0.29665 | 0.43818 | 0.21602 | 0.57095 | 0.21156 | 0.2582 | 0.29944
| Rel.Std.Dev. |  | 2.515% | 4.79% | 3.104% | 2.296% | 3.769% | 0.9353% | 1.952% | 2.578% | 1.418% | 3.634% | 1.511% | 1.347% | 1.873%
| PDM3 |  | -11.38% | -2.799% | -8.758% | 12.51% | -29.37% | -0.02055% | -0.2847% | 11.52% | -0.06599% | 3.083% | -8.153% | 25.74% | 4.854%
| Z-Score (Absolute) |  | 1.01 | 0.24 | 0.87 | 1.12 | 2.62 | 0 | 0.02 | 1.03 | 0 | 0.28 | 0.67 | 2.3 | 0.44
| Analytical results for Bi in OREAS 464 (Certified Value 4.05 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.29 | 4.07 | 4.17 | 4.18 | 3.63 | 3.897 | 4.2 | 4.4 | < 0.01 | 4.12 | 3.8264 | 3.6 | 3.9 |  | 1
| 1 | 2 | 4.14 | 4.06 | 4.11 | 4.1 | 3.58 | 3.9294 | 4.3 | 4.4 | < 0.01 | 4.2 | 3.5306 | 3.6 | 3.9 |  | #N/A
| 1 | 3 | 4.26 | 3.65 | 4.1 | 4.16 | 3.6 | 3.8733 | 4.2 | 4.5 | < 0.01 | 4.26 | 3.6509 | 3.6 | 4 |  | 16
| 1 | 4 | 4.27 | 4.01 | 4.05 | 4.13 | 3.66 | 3.899 | 4.3 | 4.4 | < 0.01 | 4.1 | 3.6172 | 3.5 | 4 |  | 4.0538
| 1 | 5 | 4.23 | 4.05 | 4.27 | 4.3 | 3.74 | 3.7816 | 4.2 | 4.3 | < 0.01 | 4.24 | 3.6808 | 3.5 | 4.1 |  | 73
| 1 | 6 | 4.31 | 4.09 | 4.06 | 4.19 | 3.66 | 3.8252 | 4.1 | 4.7 | < 0.01 | 4.15 | 3.8453 | 3.5 | 4
| Mean |  | 4.25 | 3.9883 | 4.1267 | 4.1767 | 3.645 | 3.8676 | 4.2167 | 4.45 |  | 4.1783 | 3.6919 | 3.55 | 3.9833
| Median |  | 4.265 | 4.055 | 4.105 | 4.17 | 3.645 | 3.8852 | 4.2 | 4.4 |  | 4.175 | 3.6659 | 3.55 | 4
| Std Dev. |  | 0.060332 | 0.16786 | 0.082138 | 0.068896 | 0.05648 | 0.054592 | 0.075277 | 0.13784 |  | 0.065243 | 0.12248 | 0.054772 | 0.075277
| Rel.Std.Dev. |  | 1.42% | 4.209% | 1.99% | 1.65% | 1.55% | 1.412% | 1.785% | 3.098% |  | 1.561% | 3.318% | 1.543% | 1.89%
| PDM3 |  | 4.839% | -1.616% | 1.797% | 3.03% | -10.09% | -4.594% | 4.017% | 9.773% |  | 3.071% | -8.929% | -12.43% | -1.739%
| Z-Score (Absolute) |  | 0.75 | 0.17 | 0.17 | 0.41 | 2.12 | 1.06 | 0.6 | 1.47 | NA | 0.41 | 1.9 | 2.57 | 0.51
| Analytical results for Ca in OREAS 464 (Certified Value 0.851 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.84 | 0.82 | 0.38 | 0.42 | 0.36 | 0.79497 | 0.82 | 0.92 | 0.75 | 0.9246 | 0.80297 | 0.84 | 0.877 | 0.95 | 1
| 1 | 2 | 0.84 | 0.83 | 0.37 | 0.41 | 0.36 | 0.79579 | 0.83 | 0.92 | 0.746 | 0.9413 | 0.79702 | 0.89 | 0.877 | 0.93 | #N/A
| 1 | 3 | 0.83 | 0.83 | 0.38 | 0.42 | 0.37 | 0.78263 | 0.83 | 0.91 | 0.745 | 0.9202 | 0.80992 | 0.87 | 0.8695 | 0.97 | 16
| 1 | 4 | 0.85 | 0.85 | 0.38 | 0.43 | 0.38 | 0.81991 | 0.82 | 0.92 | 0.748 | 0.8969 | 0.81383 | 0.87 | 0.8765 | 0.96 | 0.85095
| 1 | 5 | 0.83 | 0.81 | 0.37 | 0.43 | 0.36 | 0.77293 | 0.83 | 0.91 | 0.744 | 0.8886 | 0.80815 | 0.86 | 0.878 | 0.93 | 74
| 1 | 6 | 0.84 | 0.89 | 0.38 | 0.43 | 0.37 | 0.84059 | 0.82 | 0.93 | 0.734 | 0.8795 | 0.81121 | 0.87 | 0.8785 | 0.94
| Mean |  | 0.83833 | 0.83833 | 0.37667 | 0.42333 | 0.36667 | 0.80114 | 0.825 | 0.91833 | 0.7445 | 0.90852 | 0.80718 | 0.86667 | 0.87608 | 0.94667
| Median |  | 0.84 | 0.83 | 0.38 | 0.425 | 0.365 | 0.79538 | 0.825 | 0.92 | 0.7455 | 0.90855 | 0.80904 | 0.87 | 0.877 | 0.945
| Std Dev. |  | 0.0075277 | 0.028577 | 0.005164 | 0.008165 | 0.008165 | 0.024955 | 0.0054772 | 0.0075277 | 0.0055767 | 0.023847 | 0.0061628 | 0.01633 | 0.0033078 | 0.01633
| Rel.Std.Dev. |  | 0.8979% | 3.409% | 1.371% | 1.929% | 2.227% | 3.115% | 0.6639% | 0.8197% | 0.7491% | 2.625% | 0.7635% | 1.884% | 0.3776% | 1.725%
| PDM3 |  | -1.482% | -1.482% | -55.74% | -50.25% | -56.91% | -5.853% | -3.049% | 7.919% | -12.51% | 6.765% | -5.143% | 1.847% | 2.954% | 11.25%
| Z-Score (Absolute) |  | 0.12 | 0.02 | 4.61 | 4.13 | 4.71 | 0.26 | 0.02 | 0.94 | 0.84 | 0.84 | 0.2 | 0.41 | 0.51 | 1.23
| Analytical results for Cd in OREAS 464 (Certified Value 0.52 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.55 | 0.54 | 0.46 | 0.45 | 0.43 | 0.53871 | 0.7 | 0.86 | 0.51 | 0.5 | 1.5 |  |  |  | 1
| 1 | 2 | 0.53 | 0.6 | 0.55 | 0.48 | 0.46 | 0.55078 | 0.6 | 0.84 | 0.38 | 0.5 | 1 |  |  |  | 18
| 1 | 3 | 0.58 | 0.59 | 0.51 | 0.52 | 0.42 | 0.55 | 0.7 | 0.88 | 0.35 | 0.5 | 1 |  |  |  | 16
| 1 | 4 | 0.55 | 0.63 | 0.44 | 0.54 | 0.43 | 0.59346 | 0.7 | 0.8 | 0.29 | 0.5 | 1 |  |  |  | 0.51813
| 1 | 5 | 0.5 | 0.58 | 0.58 | 0.49 | 0.43 | 0.59238 | 0.5 | 0.86 | 0.42 | 0.5 | 1 |  |  |  | 75
| 1 | 6 | 0.5 | 0.68 | 0.54 | 0.52 | 0.44 | 0.56365 | 0.6 | 0.82 | 0.35 | < 0.5 | 1
| Mean |  | 0.535 | 0.60333 | 0.51333 | 0.5 | 0.435 | 0.56483 | 0.63333 | 0.84333 | 0.38333 | 0.5 | 1.0833
| Median |  | 0.54 | 0.595 | 0.525 | 0.505 | 0.43 | 0.55722 | 0.65 | 0.85 | 0.365 | 0.5 | 1
| Std Dev. |  | 0.031464 | 0.04761 | 0.054283 | 0.032863 | 0.013784 | 0.023149 | 0.08165 | 0.029439 | 0.075277 | 0 | 0.20412
| Rel.Std.Dev. |  | 5.881% | 7.891% | 10.57% | 6.573% | 3.169% | 4.098% | 12.89% | 3.491% | 19.64% | 0% | 18.84%
| PDM3 |  | 3.256% | 16.44% | -0.9257% | -3.499% | -16.04% | 9.014% | 22.23% | 62.76% | -26.02% | -3.499% | 109.1%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.21 | 0.35 | 1.04 | 0.29 | 0.97 | 3.04 | 1.5 | 0.35 | 4.59
| Analytical results for Ce in OREAS 464 (Certified Value 1.55 wt.%) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5828 | 0.83941 | 1.5021 | > 0.5 | 1.4678 | 1.6123 | 1.5599 | 1.52 |  | 1
| 1 | 2 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5715 | 0.85922 | 1.5313 | > 0.5 | 1.4714 | 1.5574 | 1.5737 | 1.54 |  | #N/A
| 1 | 3 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5691 | 0.89569 | 1.5263 | > 0.5 | 1.4926 | 1.5379 | 1.6387 | 1.55 |  | 16
| 1 | 4 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5636 | 0.83316 | 1.4983 | > 0.5 | 1.4801 | 1.5788 | 1.5558 | 1.56 |  | 1.5483
| 1 | 5 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5773 | 0.76918 | 1.4883 | > 0.5 | 1.5131 | 1.5943 | 1.5812 | 1.57 |  | 76
| 1 | 6 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5554 | 0.87641 | 1.5078 | > 0.5 | 1.4616 | 1.6522 | 1.6453 | 1.55
| Mean |  | 07: 1.5699; 08: 0.84551; 09: 1.509; 11: 1.4811; 12: 1.5888; 17: 1.5924; 22: 1.5483
| Median |  | 07: 1.5703; 08: 0.84932; 09: 1.5049; 11: 1.4758; 12: 1.5866; 17: 1.5775; 22: 1.55
| Std Dev. |  | 07: 0.0097359; 08: 0.043993; 09: 0.016682; 11: 0.019021; 12: 0.040684; 17: 0.03951; 22: 0.017224
| Rel.Std.Dev. |  | 07: 0.6201%; 08: 5.203%; 09: 1.106%; 11: 1.284%; 12: 2.561%; 17: 2.481%; 22: 1.112%
| PDM3 |  | 07: 1.399%; 08: -45.39%; 09: -2.537%; 11: -4.338%; 12: 2.619%; 17: 2.853%; 22: 0.003829%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.36 | 11.71 | 0.66 | NA | 1.12 | 0.67 | 0.73 | 0
| Analytical results for Co in OREAS 464 (Certified Value 17 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 16.7 | 16.6 | 16.4 | 17.4 | 16.7 | 16.4107 | 17.9 | 19 | 16.7 | 17.5 | 17.1681 | 25 | 15 |  | 1
| 1 | 2 | 16.3 | 16.8 | 16.4 | 17 | 16.6 | 16.0219 | 17.8 | 19 | 16.5 | 16.8 | 17.6635 | 25 | 15 |  | #N/A
| 1 | 3 | 16.7 | 14.7 | 16.9 | 17.1 | 16.2 | 16.236 | 17.6 | 18 | 16.9 | 17.7 | 17.245 | 25 | 15 |  | 16
| 1 | 4 | 17 | 15.8 | 16.6 | 17.2 | 16.6 | 16.0575 | 17.7 | 18 | 16.4 | 17 | 17.2167 | 25 | 10 |  | 16.9903
| 1 | 5 | 16.5 | 16.8 | 16.4 | 17.6 | 16.9 | 16.261 | 17.1 | 17 | 16.9 | 17.7 | 17.3671 | 25 | 20 |  | 77
| 1 | 6 | 16.9 | 16.4 | 16 | 17.2 | 16.4 | 16.5666 | 17.4 | 19 | 16.9 | 17.1 | 16.905 | 25 | 20
| Mean |  | 16.6833 | 16.1833 | 16.45 | 17.25 | 16.5667 | 16.259 | 17.5833 | 18.3333 | 16.7167 | 17.3 | 17.2609 | 25 | 15.8333
| Median |  | 16.7 | 16.5 | 16.4 | 17.2 | 16.6 | 16.2485 | 17.65 | 18.5 | 16.8 | 17.3 | 17.2309 | 25 | 15
| Std Dev. |  | 0.25626 | 0.81588 | 0.29496 | 0.21679 | 0.24221 | 0.20734 | 0.29269 | 0.8165 | 0.22286 | 0.38471 | 0.24934 | 0 | 3.7639
| Rel.Std.Dev. |  | 1.536% | 5.042% | 1.793% | 1.257% | 1.462% | 1.275% | 1.665% | 4.454% | 1.333% | 2.224% | 1.445% | 0% | 23.77%
| PDM3 |  | -1.807% | -4.749% | -3.18% | 1.529% | -2.493% | -4.304% | 3.491% | 7.905% | -1.61% | 1.823% | 1.593% | 47.14% | -6.809%
| Z-Score (Absolute) |  | 0.49 | 0.81 | 0.86 | 0.39 | 0.67 | 1.15 | 0.91 | 2.07 | 0.44 | 0.47 | 0.41 | 12.45 | 0
| Analytical results for Cr in OREAS 464 (Certified Value 335 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 309 | 310 | 143 | 159 | 140 | 322.3758 | 337 | 337 | 390 | 278 | 400 |  |  |  | 1
| 1 | 2 | 310 | 306 | 139 | 153 | 141 | 330.0602 | 343 | 329 | 394 | 285 | 400 |  |  |  | #N/A
| 1 | 3 | 304 | 307 | 143 | 158 | 143 | 330.818 | 353 | 320 | 378 | 276 | 400 |  |  |  | 16
| 1 | 4 | 314 | 314 | 139 | 159 | 145 | 322.9994 | 344 | 319 | 372 | 294 | 400 |  |  |  | 334.6808
| 1 | 5 | 303 | 279 | 139 | 164 | 139 | 326.0025 | 347 | 318 | 366 | 282 | 400 |  |  |  | 78
| 1 | 6 | 311 | 332 | 144 | 164 | 148 | 335.4223 | 357 | 334 | 358 | 287 | 400
| Mean |  | 308.5 | 308 | 141.1667 | 159.5 | 142.6667 | 327.9464 | 346.8333 | 326.1667 | 376.3333 | 283.6667 | 400
| Median |  | 309.5 | 308.5 | 141 | 159 | 142 | 328.0314 | 345.5 | 324.5 | 375 | 283.5 | 400
| Std Dev. |  | 4.2308 | 17.0997 | 2.4014 | 4.1352 | 3.3862 | 5.0574 | 7.2226 | 8.2805 | 13.8804 | 6.532 | 0
| Rel.Std.Dev. |  | 1.371% | 5.552% | 1.701% | 2.593% | 2.374% | 1.542% | 2.082% | 2.539% | 3.688% | 2.303% | 0%
| PDM3 |  | -7.823% | -7.972% | -57.82% | -52.34% | -57.37% | -2.012% | 3.631% | -2.544% | 12.45% | -15.24% | 19.52%
| Z-Score (Absolute) |  | 0 | 0.01 | 2.94 | 2.62 | 2.92 | 0.34 | 0.67 | 0.31 | 1.19 | 0.44 | 1.61
| Analytical results for Cs in OREAS 464 (Certified Value 0.29 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.28 | 0.28 | 0.3 | 0.28 | 0.26 | 0.25138 | 0.3 | < 0.5 | 0.3 | 0.13 | 0.36023 | < 0.5 | 0.3 |  | 1
| 1 | 2 | 0.29 | 0.3 | 0.28 | 0.27 | 0.26 | 0.2379 | 0.3 | < 0.5 | 0.31 | 0.14 | 0.39273 | < 0.5 | 0.2 |  | #N/A
| 1 | 3 | 0.27 | 0.3 | 0.3 | 0.28 | 0.26 | 0.27407 | 0.3 | < 0.5 | 0.32 | 0.14 | 0.41295 | < 0.5 | 0.3 |  | 16
| 1 | 4 | 0.29 | 0.28 | 0.3 | 0.29 | 0.27 | 0.25584 | 0.3 | < 0.5 | 0.3 | 0.15 | 0.35685 | < 0.5 | 0.3 |  | 0.28775
| 1 | 5 | 0.29 | 0.51 | 0.31 | 0.3 | 0.27 | 0.27302 | 0.3 | < 0.5 | 0.31 | 0.17 | 0.38254 | < 0.5 | 0.3 |  | 79
| 1 | 6 | 0.29 | 0.27 | 0.3 | 0.3 | 0.27 | 0.27025 | 0.3 | < 0.5 | 0.31 | 0.15 | 0.33117 | < 0.5 | 0.3
| Mean |  | 0.285 | 0.32333 | 0.29833 | 0.28667 | 0.265 | 0.26041 | 0.3 |  | 0.30833 | 0.14667 | 0.37275 |  | 0.28333
| Median |  | 0.29 | 0.29 | 0.3 | 0.285 | 0.265 | 0.26304 | 0.3 |  | 0.31 | 0.145 | 0.37139 |  | 0.3
| Std Dev. |  | 0.0083666 | 0.092232 | 0.0098319 | 0.012111 | 0.0054772 | 0.014499 | 0 |  | 0.0075277 | 0.013663 | 0.029152 |  | 0.040825
| Rel.Std.Dev. |  | 2.936% | 28.53% | 3.296% | 4.225% | 2.067% | 5.568% | 0% |  | 2.441% | 9.315% | 7.821% |  | 14.41%
| PDM3 |  | -0.9554% | 12.37% | 3.678% | -0.3762% | -7.906% | -9.501% | 4.257% |  | 7.154% | -49.03% | 29.54% |  | -1.535%
| Z-Score (Absolute) |  | 0.08 | 0.03 | 0.59 | 0 | 1.1 | 1.33 | 0.67 | NA | 1.1 | 7.08 | 4.35 | NA | 0.67
| Analytical results for Cu in OREAS 464 (Certified Value 92 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 92.9 | 94.2 | 89.5 | 92.3 | 97.3 | 95.0027 | 92.9 | 108 | 89.3 | 148 | 92.172 | 81.4 | 90 | 122 | 1
| 1 | 2 | 92.2 | 94.4 | 89.3 | 90.4 | 93.2 | 93.9913 | 94.1 | 111 | 89.1 | 142 | 91.9861 | 90.4 | 90 | 120 | 19
| 1 | 3 | 92.7 | 88.9 | 89.9 | 90.9 | 91.4 | 93.7953 | 93 | 117 | 89.1 | 144 | 92.2955 | 87.6 | 90 | 126 | 16
| 1 | 4 | 96 | 93.3 | 89.8 | 91 | 93.8 | 95.0528 | 93.5 | 97 | 91.7 | 138 | 92.3413 | 88.9 | 90 | 126 | 91.8599
| 1 | 5 | 92.1 | 93.7 | 90.2 | 93.4 | 94.7 | 95.2364 | 91.1 | 88 | 90 | 135 | 92.3954 | 87.4 | 90 | 124 | 80
| 1 | 6 | 93.4 | 96.5 | 87.7 | 90.8 | 93.5 | 95.1866 | 90.9 | 122 | 90.3 | 134 | 92.1176 | 91.2 | 90 | 124
| Mean |  | 93.2167 | 93.5 | 89.4 | 91.4667 | 93.9833 | 94.7109 | 92.5833 | 107.1667 | 89.9167 | 140.1667 | 92.218 | 87.8167 | 90 | 123.6667
| Median |  | 92.8 | 93.95 | 89.65 | 90.95 | 93.65 | 95.0277 | 92.95 | 109.5 | 89.65 | 140 | 92.2338 | 88.25 | 90 | 124
| Std Dev. |  | 1.4442 | 2.5116 | 0.88994 | 1.1448 | 1.9529 | 0.64196 | 1.2999 | 12.6715 | 1.0048 | 5.4559 | 0.15405 | 3.4839 | 0 | 2.3381
| Rel.Std.Dev. |  | 1.549% | 2.686% | 0.9955% | 1.252% | 2.078% | 0.6778% | 1.404% | 11.82% | 1.118% | 3.892% | 0.167% | 3.967% | 0% | 1.891%
| PDM3 |  | 1.477% | 1.785% | -2.678% | -0.4281% | 2.312% | 3.104% | 0.7875% | 16.66% | -2.115% | 52.59% | 0.3898% | -4.402% | -2.025% | 34.63%
| Z-Score (Absolute) |  | 0.01 | 0.51 | 0.92 | 0.33 | 0.39 | 0.6 | 0.01 | 4.16 | 0.77 | 13.61 | 0.12 | 1.01 | 0.75 | 8.89
| Analytical results for Dy in OREAS 464 (Certified Value 182 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 177 | 177.5 | 172.5 | 180.5 | 180.5 | 189.7844 | 139.8 | 184.9 | 105 | 179.04 | 186.7164 | 178.5 | 191 |  | 1
| 1 | 2 | 177.5 | 175.5 | 172 | 178 | 177.5 | 184.9931 | 148 | 187.4 | 102 | 181.34 | 186.618 | 178.8 | 192 |  | #N/A
| 1 | 3 | 177 | 162 | 173 | 179 | 178.5 | 186.4419 | 147.1 | 189.8 | 111 | 187.28 | 186.5737 | 184 | 198 |  | 16
| 1 | 4 | 174.5 | 171 | 172.5 | 180 | 182.5 | 187.0597 | 144.1 | 185.7 | 105 | 179.32 | 187.3401 | 176.9 | 191 |  | 181.7073
| 1 | 5 | 177 | 170.5 | 173.5 | 186.5 | 189 | 188.9754 | 138.4 | 183.8 | 107 | 188.38 | 185.8073 | 180.2 | 197 |  | 81
| 1 | 6 | 176 | 177 | 169.5 | 180.5 | 183.5 | 183.0957 | 144.8 | 188.6 | 113 | 179.99 | 186.4277 | 184.3 | 191
| Mean |  | 176.5 | 172.25 | 172.1667 | 180.75 | 181.9167 | 186.725 | 143.7 | 186.7 | 107.1667 | 182.5583 | 186.5805 | 180.45 | 193.3333
| Median |  | 177 | 173.25 | 172.5 | 180.25 | 181.5 | 186.7508 | 144.45 | 186.55 | 106 | 180.665 | 186.5959 | 179.5 | 191.5
| Std Dev. |  | 1.0954 | 5.8374 | 1.4024 | 2.9791 | 4.1523 | 2.4814 | 3.8657 | 2.2961 | 4.1191 | 4.1745 | 0.49356 | 3.0534 | 3.266
| Rel.Std.Dev. |  | 0.6206% | 3.389% | 0.8145% | 1.648% | 2.283% | 1.329% | 2.69% | 1.23% | 3.844% | 2.287% | 0.2645% | 1.692% | 1.689%
| PDM3 |  | -2.866% | -5.205% | -5.251% | -0.5269% | 0.1152% | 2.761% | -20.92% | 2.748% | -41.02% | 0.4683% | 2.682% | -0.692% | 6.398%
| Z-Score (Absolute) |  | 0.43 | 0.88 | 0.89 | 0.09 | 0.16 | 0.68 | 3.96 | 0.67 | 7.91 | 0.23 | 0.66 | 0 | 1.39
| Analytical results for Er in OREAS 464 (Certified Value 36.9 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 39.4 | 37.7 | 43.9 | 36.9 | 39.2 | 37.1888 | 29.4 | 36.9 | 34.7 | 34.66 | 68.2043 | 36.3 | 36.5 |  | 1
| 1 | 2 | 38.5 | 37.4 | 43.3 | 36.3 | 38.2 | 37.4241 | 31.6 | 36.9 | 33.9 | 35.12 | 66.8857 | 36.75 | 37 |  | #N/A
| 1 | 3 | 38.9 | 33.3 | 43.2 | 36.5 | 39.4 | 36.2596 | 31.2 | 36.6 | 35.5 | 35 | 66.0024 | 38.05 | 39.2 |  | 16
| 1 | 4 | 38.3 | 36 | 42.5 | 36.2 | 39.3 | 36.7 | 30.3 | 37.3 | 33.6 | 34.84 | 66.889 | 36.35 | 37.7 |  | 36.9362
| 1 | 5 | 38.7 | 36.5 | 44.2 | 37.9 | 40.5 | 37.2366 | 29.8 | 36.2 | 34.6 | 36.21 | 65.9028 | 36.8 | 38.2 |  | 82
| 1 | 6 | 38.8 | 37.8 | 43.4 | 36.5 | 39.5 | 36.2438 | 30.4 | 37.5 | 34.5 | 34.72 | 66.8975 | 37.6 | 37.2
| Mean |  | 38.7667 | 36.45 | 43.4167 | 36.7167 | 39.35 | 36.8422 | 30.45 | 36.9 | 34.4667 | 35.0917 | 66.7969 | 36.975 | 37.6333
| Median |  | 38.75 | 36.95 | 43.35 | 36.5 | 39.35 | 36.9444 | 30.35 | 36.9 | 34.55 | 34.92 | 66.8874 | 36.775 | 37.45
| Std Dev. |  | 0.37771 | 1.6979 | 0.59133 | 0.62743 | 0.73417 | 0.51615 | 0.82885 | 0.46904 | 0.66533 | 0.57405 | 0.82921 | 0.70338 | 0.96471
| Rel.Std.Dev. |  | 0.9743% | 4.658% | 1.362% | 1.709% | 1.866% | 1.401% | 2.722% | 1.271% | 1.93% | 1.636% | 1.241% | 1.902% | 2.563%
| PDM3 |  | 4.956% | -1.316% | 17.55% | -0.5943% | 6.535% | -0.2546% | -17.56% | -0.09796% | -6.686% | -4.994% | 80.84% | 0.1051% | 1.887%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 2.42 | 0.19 | 0.89 | 0.05 | 2.46 | 0.03 | 0.94 | 0.79 | 11.22 | 0 | 0.25
| Analytical results for Eu in OREAS 464 (Certified Value 334 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 376 | 327 | 330 | 344 | 342 | 338.8633 | 249.2 | 345.4 | 158 | 317.68 | 336.1098 | 318.8 | 302 |  | 1
| 1 | 2 | 377 | 327 | 326 | 341 | 338 | 333.4651 | 267.9 | 352.5 | 155 | 312.15 | 335.3974 | 319.9 | 300 |  | #N/A
| 1 | 3 | 379 | 297 | 331 | 343 | 341 | 325.6455 | 261 | 352.8 | 150 | 323.6 | 335.5798 | 331 | 311 |  | 16
| 1 | 4 | 377 | 314 | 325 | 341 | 349 | 322.9648 | 250 | 347.5 | 154 | 314.69 | 336.8207 | 315.4 | 309 |  | 334.4206
| 1 | 5 | 374 | 321 | 324 | 354 | 356 | 336.0578 | 238.7 | 348.9 | 153 | 325.38 | 337.7347 | 320.7 | 311 |  | 83
| 1 | 6 | 376 | 330 | 321 | 345 | 344 | 327.1213 | 255.1 | 353.3 | 159 | 311.25 | 339.152 | 330.3 | 307
| Mean |  | 376.5 | 319.3333 | 326.1667 | 344.6667 | 345 | 330.6863 | 253.65 | 350.0667 | 154.8333 | 317.4583 | 336.7991 | 322.6833 | 306.6667
| Median |  | 376.5 | 324 | 325.5 | 343.5 | 343 | 330.2932 | 252.55 | 350.7 | 154.5 | 316.185 | 336.4652 | 320.3 | 308
| Std Dev. |  | 1.6432 | 12.3396 | 3.7639 | 4.8442 | 6.5115 | 6.3432 | 10.1569 | 3.2733 | 3.3116 | 5.9145 | 1.438 | 6.4341 | 4.6762
| Rel.Std.Dev. |  | 0.4364% | 3.864% | 1.154% | 1.405% | 1.887% | 1.918% | 4.004% | 0.9351% | 2.139% | 1.863% | 0.427% | 1.994% | 1.525%
| PDM3 |  | 12.58% | -4.511% | -2.468% | 3.064% | 3.163% | -1.117% | -24.15% | 4.679% | -53.7% | -5.072% | 0.7112% | -3.51% | -8.299%
| Z-Score (Absolute) |  | 2.04 | 0.1 | 0 | 0.67 | 0.76 | 0.18 | 2.94 | 0.97 | 6.95 | 0.35 | 0.43 | 0.14 | 0.79
| Analytical results for Fe in OREAS 464 (Certified Value 35.55 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 33.6 | 35.7 | 34.5 | 37.6 | 33.8 | 38.257 | 34.13 | 37.32 | 35.4 | 36.4 | 33.192 | > 15 | 36.1085 | 36.3 | 1
| 1 | 2 | 33.7 | 35.3 | 34.3 | 37.6 | 35.4 | 38.3827 | 34.41 | 35.97 | 35.4 | 37.27 | 32.9459 | > 15 | 36.261 | 35.9 | #N/A
| 1 | 3 | 33.2 | 35.2 | 33.9 | 37.5 | 33.2 | 38.5951 | 34.65 | 36.12 | 35.4 | 36.22 | 33.2037 | > 15 | 36.048 | 37.3 | 16
| 1 | 4 | 33.6 | 34.6 | 34.7 | 38.9 | 35.4 | 38.3439 | 34.16 | 36.66 | 35.9 | 35.1 | 33.4527 | > 15 | 36.2385 | 36.8 | 35.551
| 1 | 5 | 33.3 | 34.9 | 34.6 | 38.9 | 34.7 | 37.6751 | 34.36 | 35.51 | 35.3 | 35.03 | 33.4198 | > 15 | 36.185 | 36.8 | 84
| 1 | 6 | 34.4 | 35.9 | 34.5 | 37.5 | 34.1 | 37.8 | 34.29 | 37.28 | 35.6 | 34.68 | 33.4838 | > 15 | 36.2235 | 37
| Mean |  | 33.6333 | 35.2667 | 34.4167 | 38 | 34.4333 | 38.1756 | 34.3333 | 36.4767 | 35.5 | 35.7833 | 33.283 |  | 36.1774 | 36.6833
| Median |  | 33.6 | 35.25 | 34.5 | 37.6 | 34.4 | 38.3005 | 34.325 | 36.39 | 35.4 | 35.66 | 33.3118 |  | 36.2043 | 36.8
| Std Dev. |  | 0.42269 | 0.48442 | 0.28577 | 0.69857 | 0.89144 | 0.35929 | 0.18981 | 0.73601 | 0.21909 | 1.0033 | 0.20786 |  | 0.082942 | 0.50365
| Rel.Std.Dev. |  | 1.257% | 1.374% | 0.8303% | 1.838% | 2.589% | 0.9411% | 0.5528% | 2.018% | 0.6172% | 2.804% | 0.6245% |  | 0.2293% | 1.373%
| PDM3 |  | -5.394% | -0.7997% | -3.191% | 6.889% | -3.144% | 7.383% | -3.425% | 2.604% | -0.1434% | 0.6536% | -6.38% |  | 1.762% | 3.185%
| Z-Score (Absolute) |  | 1.16 | 0.15 | 0.67 | 1.56 | 0.66 | 1.67 | 0.73 | 0.61 | 0 | 0.18 | 1.38 | NA | 0.42 | 0.74
| Analytical results for Ga in OREAS 464 (Indicative Value 97 ppm) | [15 empty] | Indicative
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 10; 08: 12; 09: 22 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS | 0
0
| 1 | 1 | 01: 109.5; 02: 83; 03: 93.4; 05: 160.5; 06: 37.5221; 07: < 0.05; 08: 103.9513; 09: FL | 1
| 1 | 2 | 01: 112.5; 02: 82.6; 03: 92.2; 05: 159.5; 06: 35.6445; 07: < 0.05; 08: 103.9887; 09: FL | 21
| 1 | 3 | 01: 102; 02: 83.6; 03: 92.5; 05: 155; 06: 39.7219; 07: < 0.05; 08: 103.2342; 09: FL | 16
| 1 | 4 | 01: 105; 02: 84.2; 03: 91.7; 05: 160; 06: 38.5406; 07: < 0.05; 08: 106.35; 09: FL | 97.0371
| 1 | 5 | 01: 108; 02: 81.7; 03: 96; 05: 161; 06: 36.2705; 07: < 0.05; 08: 106.8086; 09: FL | 27
| 1 | 6 | 01: 113; 02: 80.1; 03: 91.9; 05: 156; 06: 40.2133; 07: < 0.05; 08: 105.3179; 09: FL
| Mean |  | 01: 108.3333; 02: 82.5333; 03: 92.95; 05: 158.6667; 06: 37.9855; 08: 104.9418
| Median |  | 01: 108.75; 02: 82.8; 03: 92.35; 05: 159.75; 06: 38.0313; 08: 104.6533
| Std Dev. |  | 01: 4.2856; 02: 1.4665; 03: 1.6084; 05: 2.5232; 06: 1.8397; 08: 1.4432
| Rel.Std.Dev. |  | 01: 3.956%; 02: 1.777%; 03: 1.73%; 05: 1.59%; 06: 4.843%; 08: 1.375%
| PDM3 |  | 01: 11.64%; 02: -14.95%; 03: -4.212%; 05: 63.51%; 06: -60.85%; 08: 8.146%
| Z-Score (Absolute) |  | 01: 0.51; 02: 0.84; 03: 0.33; 05: 3.14; 06: 3.17; 07: NA; 08: 0.33; 09: NA
| Analytical results for Gd in OREAS 464 (Certified Value 670 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 690 | 679 | 661 | 610 | 689 | 683.1972 | 535.3 | 690.9 | 142 | 667.82 | 843.5711 | 660.5 | 643 |  | 1
| 1 | 2 | 693 | 668 | 656 | 603 | 680 | 675.2026 | 532.3 | 708.3 | 146 | 670.04 | 843.0187 | 659 | 640 |  | #N/A
| 1 | 3 | 687 | 595 | 658 | 617 | 680 | 657.3935 | 548.1 | 706 | 144 | 671.69 | 841.5674 | 686.3 | 660 |  | 16
| 1 | 4 | 693 | 670 | 655 | 602 | 693 | 669.2507 | 545.2 | 706.7 | 147 | 659.72 | 842.0982 | 655.6 | 669 |  | 670.037
| 1 | 5 | 699 | 703 | 663 | 636 | 715 | 690.2606 | 519.3 | 694.3 | 151 | 677.27 | 844.6831 | 670.3 | 678 |  | 85
| 1 | 6 | 685 | 692 | 646 | 608 | 693 | 661.6857 | 550.2 | 708.7 | 146 | 658.29 | 842.0767 | 690.4 | 655
| Mean |  | 691.1667 | 667.8333 | 656.5 | 612.6667 | 691.6667 | 672.8317 | 538.4 | 702.4833 | 146 | 667.4717 | 842.8359 | 670.35 | 657.5
| Median |  | 691.5 | 674.5 | 657 | 609 | 691 | 672.2266 | 540.25 | 706.35 | 146 | 668.93 | 842.5584 | 665.4 | 657.5
| Std Dev. |  | 4.9967 | 38.0916 | 5.9582 | 12.6438 | 12.8634 | 12.5875 | 11.7576 | 7.7945 | 3.0332 | 7.2789 | 1.1589 | 14.8311 | 14.7071
| Rel.Std.Dev. |  | 0.7229% | 5.704% | 0.9076% | 2.064% | 1.86% | 1.871% | 2.184% | 1.11% | 2.078% | 1.091% | 0.1375% | 2.212% | 2.237%
| PDM3 |  | 3.154% | -0.3289% | -2.02% | -8.562% | 3.228% | 0.4171% | -19.65% | 4.842% | -78.21% | -0.3829% | 25.79% | 0.04671% | -1.871%
| Z-Score (Absolute) |  | 0.67 | 0.39 | 0.45 | 2.02 | 0.69 | 0.08 | 4.27 | 1.04 | 16.99 | 0.09 | 5.59 | 0 | 0.42
| Analytical results for Ge in OREAS 464 (Indicative Value 18.4 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 11.65 | 18.45 | 16.5 | 15 | 33.7 | 0.5 | 14.6 | 34 | < 0.05 | 21.9 |  |  |  |  | 1
| 1 | 2 | 11.3 | 17.95 | 16.75 | 15 | 33.1 | 0.5 | 15.7 | 34 | < 0.05 | 21 |  |  |  |  | 15
| 1 | 3 | 11.7 | 16.05 | 16.7 | 14.9 | 31.9 | 0.4 | 15.5 | 36.8 | < 0.05 | 21.57 |  |  |  |  | 16
| 1 | 4 | 11.8 | 16.9 | 16.15 | 14.8 | 33 | 0.4 | 14.2 | 36.1 | < 0.05 | 22.41 |  |  |  |  | 18.3781
| 1 | 5 | 11.45 | 17.25 | 16.25 | 15.45 | 34 | 0.5693 | 13.5 | 35.1 | < 0.05 | 21.24 |  |  |  |  | 27
| 1 | 6 | 11.45 | 17.6 | 16 | 15.05 | 32.8 | 0.50585 | 15.2 | 35.1 | < 0.05 | 21.02
| Mean |  | 11.5583 | 17.3667 | 16.3917 | 15.0333 | 33.0833 | 0.47919 | 14.7833 | 35.1833 |  | 21.5233
| Median |  | 11.55 | 17.425 | 16.375 | 15 | 33.05 | 0.5 | 14.9 | 35.1 |  | 21.405
| Std Dev. |  | 0.18819 | 0.84063 | 0.30564 | 0.22286 | 0.73598 | 0.066691 | 0.84242 | 1.1197 |  | 0.55457
| Rel.Std.Dev. |  | 1.628% | 4.841% | 1.865% | 1.482% | 2.225% | 13.92% | 5.698% | 3.182% |  | 2.577%
| PDM3 |  | -37.11% | -5.503% | -10.81% | -18.2% | 80.02% | -97.39% | -19.56% | 91.44% |  | 17.11%
| Z-Score (Absolute) |  | 0.67 | 0.14 | 0 | 0.19 | 2.33 | 2.22 | 0.22 | 2.62 | NA | 0.72
| Analytical results for Hf in OREAS 464 (Certified Value 7.79 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 9.5 | 6.9 | 8.8 | 7.7 | 7.9 | 5.6756 | 6.9 | 1.18 | 7.81 | 10 |  |  |  |  | 1
| 1 | 2 | 9.2 | 7 | 8.9 | 7.8 | 7.6 | 6.0314 | 6.3 | 1.01 | 7.28 | 10 |  |  |  |  | #N/A
| 1 | 3 | 9.2 | 6.2 | 8.6 | 7.8 | 7.7 | 5.5216 | 6.2 | 1.08 | 7.91 | 10 |  |  |  |  | 16
| 1 | 4 | 9.3 | 6.6 | 8.7 | 7.9 | 7.8 | 5.7156 | 6.4 | 1.08 | 7.95 | 10 |  |  |  |  | 7.7917
| 1 | 5 | 8.9 | 6.1 | 9 | 8 | 8 | 5.4057 | 6.4 | 1.16 | 7.6 | 10 |  |  |  |  | 86
| 1 | 6 | 9.2 | 7 | 8.6 | 7.6 | 7.8 | 6.1802 | 6.8 | 1.14 | 7.37 | 10
| Mean |  | 9.2167 | 6.6333 | 8.7667 | 7.8 | 7.8 | 5.755 | 6.5 | 1.1083 | 7.6533 | 10
| Median |  | 9.2 | 6.75 | 8.75 | 7.8 | 7.8 | 5.6956 | 6.4 | 1.11 | 7.705 | 10
| Std Dev. |  | 0.19408 | 0.40332 | 0.1633 | 0.14142 | 0.14142 | 0.29717 | 0.28284 | 0.063377 | 0.28317 | 0
| Rel.Std.Dev. |  | 2.106% | 6.08% | 1.863% | 1.813% | 1.813% | 5.164% | 4.351% | 5.718% | 3.7% | 0%
| PDM3 |  | 18.29% | -14.87% | 12.51% | 0.1069% | 0.1069% | -26.14% | -16.58% | -85.78% | -1.775% | 28.34%
| Z-Score (Absolute) |  | 0.87 | 0.64 | 0.6 | 0.04 | 0.04 | 1.15 | 0.71 | 3.85 | 0.04 | 1.32
| Analytical results for Ho in OREAS 464 (Certified Value 21.3 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 23.2 | 20.1 | 21.8 | 20.8 | 21.7 | 20.9285 | 17 | 20.8 | 7.41 | 20.28 | 26.7814 | 21 | 21.4 |  | 1
| 1 | 2 | 23.3 | 21.4 | 21.7 | 20.7 | 21.3 | 20.5754 | 17.9 | 21.5 | 7.86 | 19.85 | 25.8772 | 21.15 | 22.4 |  | #N/A
| 1 | 3 | 23.3 | 19.8 | 21.5 | 20.8 | 21.6 | 20.5605 | 17.6 | 21.1 | 7.6 | 20.57 | 25.1493 | 21.8 | 21.8 |  | 16
| 1 | 4 | 23.1 | 21 | 21.6 | 20.6 | 21.9 | 20.4383 | 17.6 | 21.2 | 7.28 | 20.27 | 25.5899 | 20.6 | 21.8 |  | 21.3127
| 1 | 5 | 23.1 | 19.65 | 21.9 | 21.7 | 22.7 | 20.82 | 16.7 | 20.9 | 7.58 | 20.9 | 25.9525 | 21.1 | 22.9 |  | 87
| 1 | 6 | 23.4 | 21 | 21.3 | 20.9 | 22 | 20.4192 | 17.4 | 21.1 | 7.42 | 19.84 | 26.9036 | 21.55 | 21.3
| Mean |  | 23.2333 | 20.4917 | 21.6333 | 20.9167 | 21.8667 | 20.6237 | 17.3667 | 21.1 | 7.525 | 20.285 | 26.0423 | 21.2 | 21.9333
| Median |  | 23.25 | 20.55 | 21.65 | 20.8 | 21.8 | 20.5679 | 17.5 | 21.1 | 7.5 | 20.275 | 25.9149 | 21.125 | 21.8
| Std Dev. |  | 0.12111 | 0.7324 | 0.21602 | 0.39707 | 0.4761 | 0.20687 | 0.44121 | 0.24495 | 0.20256 | 0.41157 | 0.68199 | 0.42308 | 0.6121
| Rel.Std.Dev. |  | 0.5213% | 3.574% | 0.9986% | 1.898% | 2.177% | 1.003% | 2.541% | 1.161% | 2.692% | 2.029% | 2.619% | 1.996% | 2.791%
| PDM3 |  | 9.012% | -3.852% | 1.504% | -1.858% | 2.599% | -3.233% | -18.51% | -0.998% | -64.69% | -4.822% | 22.19% | -0.5288% | 2.912%
| Z-Score (Absolute) |  | 1.88 | 0.54 | 0.47 | 0.3 | 0.67 | 0.42 | 3.28 | 0 | 11.94 | 0.72 | 4.35 | 0.09 | 0.73
| Analytical results for In in OREAS 464 (Certified Value 2.11 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.17 | 2.15 | 2.46 | 2.15 | 2.17 | 2.0199 | 2.3 | 2.2 | 2.12 | 2 | 2.1 | 1.88 |  |  | 1
| 1 | 2 | 2.16 | 2.09 | 2.44 | 2.12 | 2.17 | 1.991 | 2.3 | 2.28 | 2.13 | 2.01 | 2.1 | 1.8 |  |  | #N/A
| 1 | 3 | 2.21 | 1.905 | 2.42 | 2.09 | 2.12 | 1.9906 | 2.3 | 2.22 | 2.14 | 2.01 | 2.2 | 1.94 |  |  | 16
| 1 | 4 | 2.2 | 1.99 | 2.43 | 2.18 | 2.2 | 2.0049 | 2.3 | 2.13 | 2.16 | 2.01 | 2.1 | 1.84 |  |  | 2.1115
| 1 | 5 | 2.18 | 2.05 | 2.55 | 2.32 | 2.28 | 1.9995 | 2.3 | 2.19 | 2.16 | 2.17 | 2.2 | 1.96 |  |  | 88
| 1 | 6 | 2.23 | 2.09 | 2.47 | 2.12 | 2.18 | 2.015 | 2.2 | 2.23 | 2.15 | 2.03 | 2.1 | 1.9
| Mean |  | 2.1917 | 2.0458 | 2.4617 | 2.1633 | 2.1867 | 2.0035 | 2.2833 | 2.2083 | 2.1433 | 2.0383 | 2.1333 | 1.8867
| Median |  | 2.19 | 2.07 | 2.45 | 2.135 | 2.175 | 2.0022 | 2.3 | 2.21 | 2.145 | 2.01 | 2.1 | 1.89
| Std Dev. |  | 0.026394 | 0.086857 | 0.047081 | 0.082624 | 0.052789 | 0.012177 | 0.040825 | 0.049565 | 0.01633 | 0.065243 | 0.05164 | 0.060222
| Rel.Std.Dev. |  | 1.204% | 4.246% | 1.913% | 3.819% | 2.414% | 0.6078% | 1.788% | 2.244% | 0.7619% | 3.201% | 2.421% | 3.192%
| PDM3 |  | 3.796% | -3.11% | 16.58% | 2.454% | 3.559% | -5.117% | 8.137% | 4.585% | 1.507% | -3.466% | 1.033% | -10.65%
| Z-Score (Absolute) |  | 0.44 | 0.77 | 2.68 | 0.05 | 0.4 | 1.12 | 1.2 | 0.58 | 0.04 | 1.05 | 0.04 | 2.09
| Analytical results for K in OREAS 464 (Certified Value 0.081 wt.%) | [15 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.07 | 0.09 | 0.09 | 0.07 | 0.1 | 0.1 | 0.09 | 0.089 | 0.0695 | 0.07 | 0.07 | 0.06 |  |  | 1
| 1 | 2 | 0.07 | 0.09 | 0.09 | 0.07 | 0.1 | 0.1 | 0.08 | 0.088 | 0.0737 | 0.08 | 0.07 | 0.06 |  |  | #N/A
| 1 | 3 | 0.07 | 0.1 | 0.09 | 0.07 | 0.09 | 0.1 | 0.09 | 0.089 | 0.0711 | 0.08 | 0.07 | 0.06 |  |  | 16
| 1 | 4 | 0.07 | 0.11 | 0.09 | 0.07 | 0.1 | 0.1 | 0.08 | 0.086 | 0.0644 | 0.08 | 0.07 | 0.06 |  |  | 0.081232
| 1 | 5 | 0.07 | 0.19 | 0.09 | 0.07 | 0.1 | 0.1 | 0.07 | 0.088 | 0.0695 | 0.07 | 0.07 | 0.06 |  |  | 89
| 1 | 6 | 0.07 | 0.11 | 0.09 | 0.07 | 0.1 | 0.1 | 0.09 | 0.088 | 0.0625 | 0.08 | 0.07 | 0.06
| Mean |  | 0.07 | 0.115 | 0.09 | 0.07 | 0.098333 | 0.1 | 0.083333 | 0.088 | 0.06845 | 0.076667 | 0.07 | 0.06
| Median |  | 0.07 | 0.105 | 0.09 | 0.07 | 0.1 | 0.1 | 0.085 | 0.088 | 0.0695 | 0.08 | 0.07 | 0.06
| Std Dev. |  | 0 | 0.037815 | 1.5202e-17 | 0 | 0.0040825 | 1.5202e-17 | 0.008165 | 0.0010954 | 0.0042094 | 0.005164 | 0 | 0
| Rel.Std.Dev. |  | 0% | 32.88% | 1.689e-14% | 0% | 4.152% | 1.52e-14% | 9.798% | 1.245% | 6.15% | 6.736% | 0% | 0%
| PDM3 |  | -13.83% | 41.57% | 10.79% | -13.83% | 21.05% | 23.1% | 2.587% | 8.332% | -15.74% | -5.62% | -13.83% | -26.14%
| Z-Score (Absolute) |  | 0.67 | 1.35 | 0.67 | 0.67 | 1.24 | 1.35 | 0.22 | 0.54 | 0.78 | 0.22 | 0.67 | 1.35
| Analytical results for La in OREAS 464 (Certified Value 1.06 wt.%) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | > 1 | > 1 | 0.485 | 0.497 | 0.434 | 1.2169 | 0.57946 | 1.1371 | > 0.5 | > 0.5 | 1.2141 | 1.1544 | 1.11 |  | 1
| 1 | 2 | 0.99 | > 1 | 0.473 | 0.494 | 0.436 | 1.2351 | 0.59222 | 1.1697 | > 0.5 | > 0.5 | 1.1776 | 1.1583 | 1.13 |  | #N/A
| 1 | 3 | 0.982 | > 1 | 0.478 | 0.499 | 0.442 | 1.2133 | 0.63512 | 1.1485 | > 0.5 | > 0.5 | 1.1872 | 1.2001 | 1.15 |  | 16
| 1 | 4 | > 1 | > 1 | 0.476 | 0.51 | 0.458 | 1.2224 | 0.57416 | 1.1307 | > 0.5 | > 0.5 | 1.1844 | 1.1442 | 1.14 |  | 1.0644
| 1 | 5 | 0.976 | > 1 | 0.474 | 0.522 | 0.435 | 1.2048 | 0.53198 | 1.1227 | > 0.5 | > 0.5 | 1.1948 | 1.1695 | 1.15 |  | 90
| 1 | 6 | 1 | > 1 | 0.478 | 0.51 | 0.451 | 1.2161 | 0.64198 | 1.1601 | > 0.5 | > 0.5 | 1.2408 | 1.2058 | 1.14
| Mean |  | 0.987 |  | 0.47733 | 0.50533 | 0.44267 | 1.2181 | 0.59249 | 1.1448 |  |  | 1.1998 | 1.172 | 1.1367
| Median |  | 0.986 |  | 0.477 | 0.5045 | 0.439 | 1.2165 | 0.58584 | 1.1428 |  |  | 1.191 | 1.1639 | 1.14
| Std Dev. |  | 0.010392 |  | 0.004274 | 0.010577 | 0.0098319 | 0.010148 | 0.04107 | 0.01797 |  |  | 0.023679 | 0.025342 | 0.015055
| Rel.Std.Dev. |  | 1.053% |  | 0.8954% | 2.093% | 2.221% | 0.8331% | 6.932% | 1.57% |  |  | 1.974% | 2.162% | 1.325%
| PDM3 |  | -7.273% |  | -55.16% | -52.53% | -58.41% | 14.44% | -44.34% | 7.553% |  |  | 12.72% | 10.11% | 6.787%
| Z-Score (Absolute) |  | 0.34 | NA | 2.68 | 2.55 | 2.84 | 0.72 | 2.15 | 0.38 | NA | NA | 0.63 | 0.51 | 0.34
| Analytical results for Li in OREAS 464 (Certified Value 8.26 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 2
2
| 1 | 1 | 6.8 | 6.9 | 7.6 | 8.3 | 5.1 | 8.583 | 7.3 | 9 | 7.9 | 3.1 | 9.852 | 9 | 10 | 8 | 1
| 1 | 2 | 6.9 | 6.7 | 7.6 | 8.1 | 4.5 | 8.3301 | 7.4 | 8 | 7.9 | 3 | 10.1666 | 10 | 10 | 8 | 16
| 1 | 3 | 7.2 | 6.2 | 7.9 | 8.2 | 4.7 | 7.3944 | 7.3 | 9 | 8 | 3 | 9.8194 | 9 | 10 | 8 | 16
| 1 | 4 | 7.2 | 6.6 | 7.7 | 8.3 | 4.6 | 7.441 | 7.3 | 9 | 7.8 | 3 | 10.1636 | 9 | 10 | 7.5 | 8.2553
| 1 | 5 | 7.3 | 7.2 | 7.7 | 8.3 | 4.3 | 8.0299 | 7.2 | 8 | 8.2 | 3.4 | 9.8651 | 9 | 10 | 8 | 91
| 1 | 6 | 7.5 | 6.8 | 7.8 | 8.3 | 4.1 | 8.4849 | 7.1 | 9 | 8 | 3.2 | 10.2536 | 10 | 10 | 8
| Mean |  | 7.15 | 6.7333 | 7.7167 | 8.25 | 4.55 | 8.0439 | 7.2667 | 8.6667 | 7.9667 | 3.1167 | 10.02 | 9.3333 | 10 | 7.9167
| Median |  | 7.2 | 6.75 | 7.7 | 8.3 | 4.55 | 8.18 | 7.3 | 9 | 7.95 | 3.05 | 10.0143 | 9 | 10 | 8
| Std Dev. |  | 0.25884 | 0.33267 | 0.1169 | 0.083666 | 0.34496 | 0.52011 | 0.10328 | 0.5164 | 0.13663 | 0.16021 | 0.1945 | 0.5164 | 0 | 0.20412
| Rel.Std.Dev. |  | 3.62% | 4.941% | 1.515% | 1.014% | 7.582% | 6.466% | 1.421% | 5.958% | 1.715% | 5.14% | 1.941% | 5.533% | 0% | 2.578%
| PDM3 |  | -13.39% | -18.44% | -6.525% | -0.06453% | -44.88% | -2.561% | -11.98% | 4.983% | -3.497% | -62.25% | 21.38% | 13.06% | 21.13% | -4.102%
| Z-Score (Absolute) |  | 0.7 | 1.07 | 0.2 | 0.27 | 3.02 | 0.09 | 0.6 | 0.64 | 0.02 | 4.34 | 1.85 | 1.24 | 1.83 | 0.02
| Analytical results for Lu in OREAS 464 (Certified Value 1.57 ppm) | [15 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 2.38 | 1.85 | 1.74 | 1.62 | 1.79 | 1.3708 | 1.2 | 1.41 | 1.05 | 1.38 | 2.3137 | 1.7 | 1.58 |  | 1
| 1 | 2 | 2.31 | 1.83 | 1.73 | 1.56 | 1.76 | 1.3515 | 1.2 | 1.49 | 1.13 | 1.4 | 2.2168 | 1.6 | 1.6 |  | #N/A
| 1 | 3 | 2.33 | 1.6 | 1.72 | 1.58 | 1.76 | 1.3216 | 1.2 | 1.48 | 1.19 | 1.38 | 2.2014 | 1.7 | 1.64 |  | 16
| 1 | 4 | 2.35 | 1.76 | 1.69 | 1.55 | 1.81 | 1.3831 | 1.2 | 1.49 | 1.07 | 1.4 | 2.1555 | 1.7 | 1.6 |  | 1.5722
| 1 | 5 | 2.32 | 1.67 | 1.82 | 1.61 | 1.84 | 1.2648 | 1.2 | 1.49 | 1.03 | 1.41 | 2.1625 | 1.8 | 1.66 |  | 92
| 1 | 6 | 2.33 | 1.83 | 1.73 | 1.58 | 1.78 | 1.2717 | 1.2 | 1.37 | 1.09 | 1.35 | 2.3384 | 1.6 | 1.64
| Mean |  | 2.3367 | 1.7567 | 1.7383 | 1.5833 | 1.79 | 1.3272 | 1.2 | 1.455 | 1.0933 | 1.3867 | 2.2314 | 1.6833 | 1.62
| Median |  | 2.33 | 1.795 | 1.73 | 1.58 | 1.785 | 1.3365 | 1.2 | 1.485 | 1.08 | 1.39 | 2.2091 | 1.7 | 1.62
| Std Dev. |  | 0.025033 | 0.10152 | 0.043551 | 0.027325 | 0.030984 | 0.050238 | 0 | 0.052058 | 0.058538 | 0.021602 | 0.077261 | 0.075277 | 0.030984
| Rel.Std.Dev. |  | 1.071% | 5.779% | 2.505% | 1.726% | 1.731% | 3.785% | 0% | 3.578% | 5.354% | 1.558% | 3.462% | 4.472% | 1.913%
| PDM3 |  | 48.63% | 11.74% | 10.57% | 0.7105% | 13.86% | -15.58% | -23.67% | -7.452% | -30.46% | -11.8% | 41.93% | 7.071% | 3.043%
| Z-Score (Absolute) |  | 2.84 | 0.54 | 0.4 | 0.15 | 0.67 | 1.16 | 1.67 | 0.59 | 2.09 | 0.93 | 2.43 | 0.25 | 0
| Analytical results for Mg in OREAS 464 (Certified Value 0.825 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.75 | 0.78 | 0.36 | 0.37 | 0.32 | 0.8939 | 0.78 | 0.85 | 0.863 | 0.897 | 0.77851 | 0.76 | 0.8455 | 0.85 | 1
| 1 | 2 | 0.75 | 0.77 | 0.35 | 0.37 | 0.33 | 0.84191 | 0.81 | 0.85 | 0.859 | 0.9153 | 0.77928 | 0.83 | 0.851 | 0.85 | #N/A
| 1 | 3 | 0.74 | 0.77 | 0.35 | 0.37 | 0.33 | 0.84424 | 0.81 | 0.85 | 0.864 | 0.8939 | 0.78226 | 0.8 | 0.839 | 0.89 | 16
| 1 | 4 | 0.77 | 0.77 | 0.35 | 0.38 | 0.34 | 0.86157 | 0.81 | 0.85 | 0.852 | 0.8702 | 0.79186 | 0.8 | 0.847 | 0.87 | 0.82488
| 1 | 5 | 0.74 | 0.75 | 0.35 | 0.39 | 0.33 | 0.85368 | 0.8 | 0.84 | 0.862 | 0.8596 | 0.78656 | 0.78 | 0.8495 | 0.85 | 93
| 1 | 6 | 0.76 | 0.79 | 0.35 | 0.38 | 0.33 | 0.87992 | 0.81 | 0.86 | 0.867 | 0.8556 | 0.78868 | 0.82 | 0.849 | 0.86
| Mean |  | 0.75167 | 0.77167 | 0.35167 | 0.37667 | 0.33 | 0.86254 | 0.80333 | 0.85 | 0.86117 | 0.88193 | 0.78453 | 0.79833 | 0.84683 | 0.86167
| Median |  | 0.75 | 0.77 | 0.35 | 0.375 | 0.33 | 0.85762 | 0.81 | 0.85 | 0.8625 | 0.88205 | 0.78441 | 0.8 | 0.848 | 0.855
| Std Dev. |  | 0.01169 | 0.013292 | 0.0040825 | 0.008165 | 0.0063246 | 0.02062 | 0.012111 | 0.0063246 | 0.0051929 | 0.023721 | 0.0053701 | 0.025626 | 0.0042973 | 0.016021
| Rel.Std.Dev. |  | 1.555% | 1.722% | 1.161% | 2.168% | 1.917% | 2.391% | 1.508% | 0.7441% | 0.603% | 2.69% | 0.6845% | 3.21% | 0.5075% | 1.859%
| PDM3 |  | -8.875% | -6.451% | -57.37% | -54.34% | -59.99% | 4.565% | -2.612% | 3.046% | 4.399% | 6.917% | -4.892% | -3.218% | 2.662% | 4.46%
| Z-Score (Absolute) |  | 0.61 | 0.36 | 5.53 | 5.22 | 5.8 | 0.76 | 0.03 | 0.61 | 0.74 | 1 | 0.2 | 0.03 | 0.57 | 0.75
| Analytical results for Mn in OREAS 464 (Certified Value 0.27 wt.%) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.239 | 0.267 | 0.0964 | 0.0976 | 0.086 | 0.27101 | 0.27551 | 0.2746 | 0.277 | 0.2903 | 0.242 | 0.2815 | 0.282 |  | 1
| 1 | 2 | 0.238 | 0.268 | 0.0933 | 0.0957 | 0.0851 | 0.26505 | 0.27892 | 0.2646 | 0.277 | 0.3021 | 0.2532 | 0.2815 | 0.279 |  | #N/A
| 1 | 3 | 0.236 | 0.269 | 0.0958 | 0.0969 | 0.0903 | 0.27094 | 0.2787 | 0.2667 | 0.279 | 0.2902 | 0.2322 | 0.2805 | 0.29 |  | 16
| 1 | 4 | 0.242 | 0.277 | 0.093 | 0.0972 | 0.0911 | 0.26964 | 0.2738 | 0.2713 | 0.272 | 0.2824 | 0.2479 | 0.281 | 0.286 |  | 0.26998
| 1 | 5 | 0.232 | 0.277 | 0.093 | 0.102 | 0.0854 | 0.26759 | 0.26937 | 0.2602 | 0.277 | 0.2798 | 0.2456 | 0.281 | 0.283 |  | 94
| 1 | 6 | 0.24 | 0.291 | 0.0939 | 0.102 | 0.0913 | 0.27426 | 0.2675 | 0.2756 | 0.28 | 0.2772 | 0.251 | 0.281 | 0.285
| Mean |  | 0.23783 | 0.27483 | 0.094233 | 0.098567 | 0.0882 | 0.26975 | 0.27397 | 0.26883 | 0.277 | 0.287 | 0.24532 | 0.28108 | 0.28417
| Median |  | 0.2385 | 0.273 | 0.0936 | 0.0974 | 0.08815 | 0.27029 | 0.27465 | 0.269 | 0.277 | 0.2863 | 0.24675 | 0.281 | 0.284
| Std Dev. |  | 0.0034881 | 0.0090866 | 0.0014949 | 0.002734 | 0.0029907 | 0.0031648 | 0.0047391 | 0.0060355 | 0.0027568 | 0.0091453 | 0.0075399 | 0.00037639 | 0.0037639
| Rel.Std.Dev. |  | 1.467% | 3.306% | 1.586% | 2.774% | 3.391% | 1.173% | 1.73% | 2.245% | 0.9952% | 3.187% | 3.074% | 0.1339% | 1.325%
| PDM3 |  | -11.91% | 1.798% | -65.1% | -63.49% | -67.33% | -0.08555% | 1.477% | -0.424% | 2.601% | 6.305% | -9.135% | 4.113% | 5.255%
| Z-Score (Absolute) |  | 1.49 | 0.24 | 8.21 | 8 | 8.49 | 0 | 0.2 | 0.04 | 0.34 | 0.81 | 1.14 | 0.53 | 0.67
| Analytical results for Mo in OREAS 464 (Certified Value 77 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
1
| 1 | 1 | 75.8 | 77.3 | 82.5 | 81.6 | 81.8 | 73.4024 | 75.7 | 81 | 80.7 | 77.7 | 73.5348 | 66.5 |  |  | 1
| 1 | 2 | 75.4 | 76.8 | 82.7 | 79.6 | 81.2 | 73.3387 | 77 | 83 | 79.4 | 73.8 | 74.8474 | 68 |  |  | #N/A
| 1 | 3 | 75.5 | 71.4 | 82 | 80 | 80.2 | 72.8595 | 75.6 | 82 | 80.9 | 77.2 | 73.4083 | 69 |  |  | 16
| 1 | 4 | 76.5 | 73.8 | 85.3 | 81.6 | 83.2 | 73.8951 | 75.4 | 79 | 79.7 | 73.9 | 73.8087 | 68.5 |  |  | 77.135
| 1 | 5 | 74.6 | 79.8 | 84.6 | 84 | 84.1 | 73.2376 | 73.4 | 78 | 80.3 | 77.9 | 73.1467 | 69.5 |  |  | 95
| 1 | 6 | 76.9 | 76.7 | 81.1 | 81.5 | 81.5 | 73.7259 | 73.5 | 80 | 80.1 | 76.1 | 73.2136 | 69.5
| Mean |  | 75.7833 | 75.9667 | 83.0333 | 81.3833 | 82 | 73.4099 | 75.1 | 80.5 | 80.1833 | 76.1 | 73.6599 | 68.5
| Median |  | 75.65 | 76.75 | 82.6 | 81.55 | 81.65 | 73.3706 | 75.5 | 80.5 | 80.2 | 76.65 | 73.4716 | 68.75
| Std Dev. |  | 0.82321 | 2.9412 | 1.5996 | 1.5523 | 1.4156 | 0.3669 | 1.3971 | 1.8708 | 0.57417 | 1.8515 | 0.62828 | 1.1402
| Rel.Std.Dev. |  | 1.086% | 3.872% | 1.926% | 1.907% | 1.726% | 0.4998% | 1.86% | 2.324% | 0.7161% | 2.433% | 0.8529% | 1.664%
| PDM3 |  | -1.752% | -1.515% | 7.647% | 5.508% | 6.307% | -4.829% | -2.638% | 4.363% | 3.952% | -1.342% | -4.505% | -11.19%
| Z-Score (Absolute) |  | 0.05 | 0.01 | 1.39 | 1.07 | 1.19 | 0.52 | 0.19 | 0.89 | 0.83 | 0.01 | 0.47 | 1.5
| Analytical results for Na in OREAS 464 (Certified Value 0.149 wt.%) | [15 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.13 | 0.14 | 0.14 | 0.16 | 0.14 | 0.14 | 0.15 | 0.158 | 0.142 | 0.15008 | 0.17 | 0.14 | 0.16 |  | 1
| 1 | 2 | 0.13 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.15 | 0.157 | 0.1452 | 0.15615 | 0.17 | 0.14 | 0.15 |  | 17
| 1 | 3 | 0.13 | 0.16 | 0.14 | 0.16 | 0.13 | 0.15 | 0.14 | 0.159 | 0.1384 | 0.15354 | 0.17 | 0.14 | 0.16 |  | 16
| 1 | 4 | 0.13 | 0.16 | 0.14 | 0.16 | 0.14 | 0.15 | 0.15 | 0.155 | 0.1361 | 0.15589 | 0.17 | 0.14 | 0.16 |  | 0.14862
| 1 | 5 | 0.13 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.14 | 0.158 | 0.1357 | 0.15531 | 0.18 | 0.14 | 0.15 |  | 96
| 1 | 6 | 0.14 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.14 | 0.159 | 0.1329 | 0.15521 | 0.17 | 0.14 | 0.15
| Mean |  | 0.13167 | 0.15167 | 0.14 | 0.16 | 0.13833 | 0.14833 | 0.145 | 0.15767 | 0.13838 | 0.15436 | 0.17167 | 0.14 | 0.155
| Median |  | 0.13 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.145 | 0.158 | 0.13725 | 0.15526 | 0.17 | 0.14 | 0.155
| Std Dev. |  | 0.0040825 | 0.0075277 | 0 | 0 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0015055 | 0.0045173 | 0.0022864 | 0.0040825 | 0 | 0.0054772
| Rel.Std.Dev. |  | 3.101% | 4.963% | 0% | 0% | 2.951% | 2.752% | 3.777% | 0.9549% | 3.264% | 1.481% | 2.378% | 0% | 3.534%
| PDM3 |  | -11.41% | 2.049% | -5.801% | 7.656% | -6.922% | -0.1939% | -2.437% | 6.086% | -6.889% | 3.863% | 15.51% | -5.801% | 4.292%
| Z-Score (Absolute) |  | 1.35 | 0.27 | 0.67 | 0.94 | 0.81 | 0 | 0.27 | 0.76 | 0.81 | 0.49 | 1.89 | 0.67 | 0.54
| Analytical results for Nb in OREAS 464 (Indicative Value 1598 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 1622.7466 | 1633.8 | 1356 | > 500 | 1860 |  |  |  |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 1679.193 | 1653.4 | 1063 | > 500 | 1880 |  |  |  |  | 21
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 1684.3296 | 1629.3 | 1167 | > 500 | 1940 |  |  |  |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 1643.3773 | 1646.1 | 1224 | > 500 | 1970 |  |  |  |  | 1597.7592
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 1630.7065 | 1612.3 | 939 | > 500 | 1950 |  |  |  |  | 27
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 1693.5686 | 1598.4 | 1300 | > 500 | 1970
| Mean |  | 07: 1658.987; 08: 1628.8833; 09: 1174.8333; 22: 1928.3333
| Median |  | 07: 1661.2852; 08: 1631.55; 09: 1195.5; 22: 1945
| Std Dev. |  | 07: 30.3421; 08: 20.6207; 09: 154.3314; 22: 47.0815
| Rel.Std.Dev. |  | 07: 1.829%; 08: 1.266%; 09: 13.14%; 22: 2.442%
| PDM3 |  | 07: 3.832%; 08: 1.948%; 09: -26.47%; 22: 20.69%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.07 | 0.07 | 2.11 | NA | 1.28
| Analytical results for Nd in OREAS 464 (Certified Value 0.959 wt.%) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | > 1 | > 1 | 0.951 | > 1 | 0.941 | 1.0498 | 0.60099 | 0.9651 | > 0.5 | 0.92443 | 0.92704 | 0.97225 | 0.947 |  | 1
| 1 | 2 | 0.991 | > 1 | 0.946 | > 1 | 0.932 | 1.0482 | 0.65359 | 0.9773 | > 0.5 | 0.92877 | 0.91585 | 0.9743 | 0.947 |  | #N/A
| 1 | 3 | 1 | 0.918 | 0.966 | > 1 | 0.924 | 1.0415 | 0.65278 | 0.9738 | > 0.5 | 0.97265 | 0.91595 | 1.0126 | 0.964 |  | 16
| 1 | 4 | 0.994 | 0.975 | 0.959 | > 1 | 0.951 | 1.0411 | 0.60996 | 0.9752 | > 0.5 | 0.93701 | 0.92544 | 0.967 | 0.97 |  | 0.95892
| 1 | 5 | 0.983 | 0.968 | 0.944 | > 1 | 0.983 | 1.0471 | 0.57539 | 0.9632 | > 0.5 | 0.95021 | 0.92112 | 0.9827 | 0.968 |  | 97
| 1 | 6 | 0.985 | 1 | 0.938 | > 1 | 0.954 | 1.0506 | 0.64421 | 0.9628 | > 0.5 | 0.92692 | 0.9271 | 1.009 | 0.954
| Mean |  | 0.9906 | 0.96525 | 0.95067 |  | 0.9475 | 1.0464 | 0.62282 | 0.96957 |  | 0.94 | 0.92208 | 0.98631 | 0.95833
| Median |  | 0.991 | 0.9715 | 0.9485 |  | 0.946 | 1.0477 | 0.62709 | 0.96945 |  | 0.93289 | 0.92328 | 0.9785 | 0.959
| Std Dev. |  | 0.0068775 | 0.034364 | 0.010309 |  | 0.020734 | 0.0041104 | 0.032229 | 0.0065686 |  | 0.018542 | 0.0052625 | 0.019666 | 0.010367
| Rel.Std.Dev. |  | 0.6943% | 3.56% | 1.084% |  | 2.188% | 0.3928% | 5.175% | 0.6775% |  | 1.973% | 0.5707% | 1.994% | 1.082%
| PDM3 |  | 3.303% | 0.6598% | -0.861% |  | -1.191% | 9.122% | -35.05% | 1.11% |  | -1.974% | -3.842% | 2.856% | -0.06148%
| Z-Score (Absolute) |  | 1.19 | 0.25 | 0.28 | NA | 0.4 | 3.24 | 12.34 | 0.41 | NA | 0.67 | 1.33 | 1.03 | 0
| Analytical results for Ni in OREAS 464 (Certified Value 89 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
1
| 1 | 1 | 362 | 85.7 | 84.5 | 88.1 | 83.4 | 83.804 | 95 | 92.8 | 95 | 80 | 100 |  |  |  | 1
| 1 | 2 | 357 | 84.5 | 83.6 | 84.6 | 82 | 85.0529 | 94 | 91.3 | 96 | 80 | 108 |  |  |  | #N/A
| 1 | 3 | 357 | 78.1 | 86.1 | 86.7 | 79.6 | 84.0586 | 89 | 92.9 | 92 | 85 | 100 |  |  |  | 16
| 1 | 4 | 365 | 83.9 | 84.6 | 85.4 | 83 | 84.6677 | 99 | 92.8 | 92 | 80 | 104 |  |  |  | 88.5748
| 1 | 5 | 358 | 86.2 | 83.8 | 89.7 | 84.8 | 85.676 | 89 | 91 | 90 | 85 | 102 |  |  |  | 98
| 1 | 6 | 369 | 85.7 | 84.4 | 86.6 | 87.5 | 84.8258 | 87 | 92 | 88 | 80 | 104
| Mean |  | 361.3333 | 84.0167 | 84.5 | 86.85 | 83.3833 | 84.6809 | 92.1667 | 92.1333 | 92.1667 | 81.6667 | 103
| Median |  | 360 | 85.1 | 84.45 | 86.65 | 83.2 | 84.7468 | 91.5 | 92.4 | 92 | 80 | 103
| Std Dev. |  | 4.9261 | 3.0229 | 0.88091 | 1.8404 | 2.6581 | 0.67912 | 4.5789 | 0.83347 | 2.9944 | 2.582 | 3.0332
| Rel.Std.Dev. |  | 1.363% | 3.598% | 1.042% | 2.119% | 3.188% | 0.802% | 4.968% | 0.9046% | 3.249% | 3.162% | 2.945%
| PDM3 |  | 307.9% | -5.146% | -4.6% | -1.947% | -5.861% | -4.396% | 4.055% | 4.018% | 4.055% | -7.799% | 16.29%
| Z-Score (Absolute) |  | 35.71 | 0.21 | 0.31 | 0 | 0.45 | 0.28 | 0.69 | 0.69 | 0.69 | 0.67 | 2.1
| Analytical results for P in OREAS 464 (Certified Value 1.36 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | > 1 | > 1 | 0.614 | 0.654 | 0.567 | 1.297 | 1.45 | 1.37 | 1.4155 | 1.2593 | 1.3 | 1.32 | 1.35 |  | 1
| 1 | 2 | > 1 | > 1 | 0.598 | 0.646 | 0.569 | 1.319 | 1.45 | 1.36 | 1.4482 | 1.2714 | 1.3 | 1.3185 | 1.49 |  | #N/A
| 1 | 3 | > 1 | > 1 | 0.605 | 0.655 | 0.584 | 1.321 | 1.46 | 1.36 | 1.395 | 1.2731 | 1.3 | 1.334 | 1.59 |  | 16
| 1 | 4 | > 1 | > 1 | 0.602 | 0.663 | 0.593 | 1.304 | 1.42 | 1.33 | 1.383 | 1.2943 | 1.3 | 1.321 | 1.46 |  | 1.3595
| 1 | 5 | > 1 | > 1 | 0.601 | 0.678 | 0.568 | 1.313 | 1.42 | 1.36 | 1.362 | 1.2873 | 1.3 | 1.306 | 1.42 |  | 99
| 1 | 6 | > 1 | > 1 | 0.605 | 0.672 | 0.586 | 1.311 | 1.48 | 1.37 | 1.3581 | 1.2981 | 1.4 | 1.3175 | 1.53
| Mean |  |  |  | 0.60417 | 0.66133 | 0.57783 | 1.3108 | 1.4467 | 1.3583 | 1.3936 | 1.2806 | 1.3167 | 1.3195 | 1.4733
| Median |  |  |  | 0.6035 | 0.659 | 0.5765 | 1.312 | 1.45 | 1.36 | 1.389 | 1.2802 | 1.3 | 1.3193 | 1.475
| Std Dev. |  |  |  | 0.0054924 | 0.012028 | 0.011197 | 0.0090866 | 0.023381 | 0.01472 | 0.034158 | 0.01505 | 0.040825 | 0.0089443 | 0.084063
| Rel.Std.Dev. |  |  |  | 0.9091% | 1.819% | 1.938% | 0.6932% | 1.616% | 1.084% | 2.451% | 1.175% | 3.101% | 0.6779% | 5.706%
| PDM3 |  |  |  | -55.56% | -51.36% | -57.5% | -3.582% | 6.41% | -0.08781% | 2.509% | -5.807% | -3.153% | -2.944% | 8.371%
| Z-Score (Absolute) |  | NA | NA | 6.24 | 5.74 | 6.47 | 0.05 | 1.14 | 0.37 | 0.67 | 0.32 | 0 | 0.02 | 1.17
| Analytical results for Pb in OREAS 464 (Certified Value 205 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 206 | 219 | 215 | 212 | 205 | 186.8226 | 175 | 190 | 146.2 | 255.535 | 260 | 173 |  |  | 1
| 1 | 2 | 199 | 211 | 226 | 227 | 210 | 183.1645 | 181 | 191 | 144.5 | 255.815 | 260 | 171 |  |  | #N/A
| 1 | 3 | 200 | 227 | 216 | 217 | 203 | 182.5035 | 177 | 199 | 152.2 | 255.201 | 260 | 174 |  |  | 16
| 1 | 4 | 205 | 208 | 229 | 231 | 207 | 185.4246 | 181 | 188 | 151.2 | 255.742 | 270 | 174 |  |  | 204.9508
| 1 | 5 | 199.5 | 206 | 230 | 220 | 213 | 186.6728 | 178 | 193 | 151.7 | 254.587 | 260 | 176 |  |  | 100
| 1 | 6 | 209 | 217 | 218 | 210 | 203 | 189.5934 | 173 | 194 | 145.6 | 256.494 | 250 | 171
| Mean |  | 203.0833 | 214.6667 | 222.3333 | 219.5 | 206.8333 | 185.6969 | 177.5 | 192.5 | 148.5667 | 255.5623 | 260 | 173.1667
| Median |  | 202.5 | 214 | 222 | 218.5 | 206 | 186.0487 | 177.5 | 192 | 148.7 | 255.6385 | 260 | 173.5
| Std Dev. |  | 4.1523 | 7.8655 | 6.7725 | 8.2644 | 4.0208 | 2.611 | 3.2094 | 3.8341 | 3.4898 | 0.6398 | 6.3246 | 1.9408
| Rel.Std.Dev. |  | 2.045% | 3.664% | 3.046% | 3.765% | 1.944% | 1.406% | 1.808% | 1.992% | 2.349% | 0.2503% | 2.433% | 1.121%
| PDM3 |  | -0.9112% | 4.741% | 8.481% | 7.099% | 0.9185% | -9.394% | -13.39% | -6.075% | -27.51% | 24.69% | 26.86% | -15.51%
| Z-Score (Absolute) |  | 0.07 | 0.36 | 0.64 | 0.54 | 0.07 | 0.71 | 1.01 | 0.46 | 2.08 | 1.86 | 2.03 | 1.17
| Analytical results for Pr in OREAS 464 (Certified Value 2601 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2766.0865 | 1499.4 | 2526.18 | 993 | 2614.19 | 2738.0649 | 2513.1 | 2460 |  | 1
| 1 | 2 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2705.2712 | 1631.1 | 2592.95 | 997 | 2677.05 | 2640.5558 | 2535.7 | 2470 |  | #N/A
| 1 | 3 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2603.8712 | 1628.9 | 2574.9 | 1050 | 2601.65 | 2604.5449 | 2621.6 | 2490 |  | 16
| 1 | 4 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2659.473 | 1517.4 | 2551.77 | 1000 | 2615.49 | 2670.0232 | 2514.4 | 2490 |  | 2600.6204
| 1 | 5 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2714.0279 | 1421.4 | 2519.2 | 1040 | 2702.63 | 2691.2284 | 2540.8 | 2500 |  | 101
| 1 | 6 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2641.0841 | 1631.5 | 2561.15 | 1060 | 2592.96 | 2786.2845 | 2636.1 | 2500
| Mean | [6 empty] | 2681.6357 | 1554.95 | 2554.3583 | 1023.3333 | 2633.995 | 2688.4503 | 2560.2833 | 2485
| Median | [6 empty] | 2682.3721 | 1573.15 | 2556.46 | 1020 | 2614.84 | 2680.6258 | 2538.25 | 2490
| Std Dev. | [6 empty] | 58.1769 | 88.8355 | 28.2639 | 29.9711 | 44.7867 | 65.9006 | 54.4489 | 16.4317
| Rel.Std.Dev. | [6 empty] | 2.169% | 5.713% | 1.106% | 2.929% | 1.7% | 2.451% | 2.127% | 0.6612%
| PDM3 | [6 empty] | 3.115% | -40.21% | -1.779% | -60.65% | 1.283% | 3.377% | -1.551% | -4.446%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.83 | 6.73 | 0.02 | 10.29 | 0.51 | 0.88 | 0.02 | 0.49
| Analytical results for Rb in OREAS 464 (Certified Value 4.22 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.7 | 4.5 | 4.3 | 4.1 | 3.9 | 4.148 | 4 | 4.7 | 3.46 | 4.17 | 5.0242 | 4 | 4.6 |  | 1
| 1 | 2 | 4.8 | 4.6 | 4.5 | 4.1 | 3.9 | 4.0603 | 4.1 | 4.9 | 3.29 | 4 | 4.86 | 4 | 4.6 |  | #N/A
| 1 | 3 | 4.4 | 4.1 | 4.3 | 4 | 3.9 | 4.2091 | 4 | 4.7 | 3.42 | 4.06 | 5.2089 | 4 | 4.6 |  | 16
| 1 | 4 | 4.9 | 4.2 | 4.4 | 4.1 | 3.9 | 4.386 | 4 | 4.9 | 3.32 | 4.17 | 5.3312 | 4 | 4.4 |  | 4.2208
| 1 | 5 | 4.8 | 4.7 | 4.5 | 4.2 | 4.1 | 4.0062 | 3.9 | 4.6 | 3.47 | 4.39 | 5.0678 | 4 | 4.6 |  | 102
| 1 | 6 | 4.9 | 4.4 | 4.2 | 4.1 | 3.9 | 3.8792 | 4 | 4.8 | 3.43 | 4.21 | 5.0192 | 4 | 4.6
| Mean |  | 4.75 | 4.4167 | 4.3667 | 4.1 | 3.9333 | 4.1148 | 4 | 4.7667 | 3.3983 | 4.1667 | 5.0852 | 4 | 4.5667
| Median |  | 4.8 | 4.45 | 4.35 | 4.1 | 3.9 | 4.1041 | 4 | 4.75 | 3.425 | 4.17 | 5.046 | 4 | 4.6
| Std Dev. |  | 0.18708 | 0.23166 | 0.12111 | 0.063246 | 0.08165 | 0.1754 | 0.063246 | 0.12111 | 0.075211 | 0.13486 | 0.16426 | 0 | 0.08165
| Rel.Std.Dev. |  | 3.939% | 5.245% | 2.773% | 1.543% | 2.076% | 4.263% | 1.581% | 2.541% | 2.213% | 3.237% | 3.23% | 0% | 1.788%
| PDM3 |  | 12.54% | 4.64% | 3.455% | -2.862% | -6.811% | -2.511% | -5.232% | 12.93% | -19.49% | -1.283% | 20.48% | -5.232% | 8.194%
| Z-Score (Absolute) |  | 1.89 | 0.72 | 0.58 | 0.19 | 0.67 | 0.15 | 0.48 | 1.73 | 2.22 | 0 | 2.65 | 0.48 | 1.16
| Analytical results for Re in OREAS 464 (Certified Value < 0.03 ppm) | [15 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.049 | 0.005 | 0.017 | 0.012 | 0.02 | < 0.002 | 0.02 | 0.005 | 0.008 | < 0.05 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.051 | 0.005 | 0.017 | 0.01 | 0.021 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 18
| 1 | 3 | 0.051 | 0.004 | 0.017 | 0.012 | 0.021 | < 0.002 | < 0.02 | 0.004 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.051 | 0.008 | 0.017 | 0.014 | 0.02 | < 0.002 | < 0.02 | 0.004 | 0.003 | < 0.05 | < 0.1 |  |  |  | < 0.03
| 1 | 5 | 0.048 | 0.002 | 0.013 | 0.017 | 0.021 | < 0.002 | < 0.02 | 0.004 | 0.003 | < 0.05 | < 0.1 |  |  |  | 103
| 1 | 6 | 0.047 | 0.015 | 0.018 | 0.011 | 0.02 | < 0.002 | < 0.02 | 0.004 | 0.003 | < 0.05 | < 0.1
| Mean |  | 0.0495 | 0.0065 | 0.0165 | 0.012667 | 0.0205 |  | 0.02 | 0.004 | 0.00425
| Median |  | 0.05 | 0.005 | 0.017 | 0.012 | 0.0205 |  | 0.02 | 0.004 | 0.003
| Std Dev. |  | 01: 0.0017607; 02: 0.0045935; 03: 0.0017607; 05: 0.0025033; 06: 0.00054772; 09: 0.00063246; 10: 0.0025
| Rel.Std.Dev. |  | 01: 3.557%; 02: 70.67%; 03: 10.67%; 05: 19.76%; 06: 2.672%; 09: 15.81%; 10: 58.82%
PDM3
| Z-Score (Absolute) |  | 3.33 | 0.81 | 0.22 | 0.22 | 0.54 | NA | 0.49 | 1.05 | 1.15 | NA | NA
| Analytical results for S in OREAS 464 (Certified Value 983 ppm) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 960 | 963 | 1100 | 1050 | 1000 |  |  | 1
| 1 | 2 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 920 | 1095 | 1100 | 1050 | 1000 |  |  | #N/A
| 1 | 3 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 940 | 934 | 1100 | 1050 | 1050 |  |  | 16
| 1 | 4 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 920 | 949 | 1200 | 1050 | 1000 |  |  | 983.4111
| 1 | 5 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 940 | 1277 | 1100 | 1050 | 1050 |  |  | 104
| 1 | 6 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 930 | 972 | 1100 | 1050 | 1000
| Mean |  | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 935 | 1031.6667 | 1116.6667 | 1050 | 1016.6667
| Median |  | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 935 | 967.5 | 1100 | 1050 | 1000
| Std Dev. |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.1658 | 133.3022 | 40.8248 | 0 | 25.8199
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 1.622% | 12.92% | 3.656% | 0% | 2.54%
| PDM3 |  | -8.482% | 11.86% | -8.482% | 1.687% | 1.687% | -8.482% | -8.482% | -4.923% | 4.907% | 13.55% | 6.771% | 3.382%
| Z-Score (Absolute) |  | 0.82 | 0.98 | 0.82 | 0.08 | 0.08 | 0.82 | 0.82 | 0.51 | 0.08 | 1.13 | 0.53 | 0.23
| Analytical results for Sb in OREAS 464 (Certified Value 1.87 ppm) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.77 | 1.92 | 2.07 | 1.93 | 1.86 | 1.617 | 1.9 | 1.8 | 1.75 | 1.47 | 1.8 |  |  |  | 1
| 1 | 2 | 1.79 | 1.86 | 1.97 | 1.87 | 1.83 | 1.6035 | 1.9 | 1.8 | 1.74 | 1.39 | 1.8 |  |  |  | #N/A
| 1 | 3 | 1.83 | 1.69 | 2.02 | 1.94 | 1.83 | 1.807 | 1.9 | 2.1 | 1.8 | 1.46 | 1.8 |  |  |  | 16
| 1 | 4 | 1.83 | 1.87 | 2.06 | 1.95 | 1.87 | 1.6575 | 1.9 | 1.9 | 1.82 | 1.47 | 1.8 |  |  |  | 1.868
| 1 | 5 | 1.8 | 1.99 | 2.17 | 2.03 | 1.94 | 1.7175 | 1.9 | 1.9 | 1.83 | 1.49 | 1.6 |  |  |  | 105
| 1 | 6 | 1.98 | 1.9 | 2 | 1.94 | 1.86 | 1.7611 | 1.8 | 2 | 1.79 | 1.43 | 1.8
| Mean |  | 1.8333 | 1.8717 | 2.0483 | 1.9433 | 1.865 | 1.6939 | 1.8833 | 1.9167 | 1.7883 | 1.4517 | 1.7667
| Median |  | 1.815 | 1.885 | 2.04 | 1.94 | 1.86 | 1.6875 | 1.9 | 1.9 | 1.795 | 1.465 | 1.8
| Std Dev. |  | 0.075542 | 0.10028 | 0.070261 | 0.051251 | 0.040373 | 0.081614 | 0.040825 | 0.1169 | 0.03656 | 0.036009 | 0.08165
| Rel.Std.Dev. |  | 4.12% | 5.358% | 3.43% | 2.637% | 2.165% | 4.818% | 2.168% | 6.099% | 2.044% | 2.481% | 4.622%
| PDM3 |  | -1.857% | 0.1949% | 9.652% | 4.031% | -0.162% | -9.32% | 0.8195% | 2.604% | -4.266% | -22.29% | -5.426%
| Z-Score (Absolute) |  | 0.33 | 0.45 | 1.9 | 0.81 | 0 | 1.77 | 0.19 | 0.54 | 0.8 | 4.29 | 0.67
| Analytical results for Sc in OREAS 464 (Certified Value 141 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 134.5 | 147 | 141.5 | 150 | 143 | 132.49 | 120 | 140 | 83.6 | 149.9738 | 140 | 143 |  |  | 1
| 1 | 2 | 135 | 141.5 | 142 | 146.5 | 140.5 | 136.0164 | 126 | 135 | 84 | 151.1792 | 139 | 143 |  |  | 19
| 1 | 3 | 134 | 132 | 147 | 148.5 | 137 | 130.3956 | 126 | 138 | 85.8 | 148.6975 | 139 | 150 |  |  | 16
| 1 | 4 | 136.5 | 140 | 143.5 | 150.5 | 142 | 136.9 | 118 | 138 | 89 | 158.299 | 139 | 147 |  |  | 141.3799
| 1 | 5 | 133 | 134.5 | 144 | 152.5 | 146 | 134.9671 | 113 | 134 | 89 | 148.4705 | 139 | 146 |  |  | 106
| 1 | 6 | 137 | 141 | 142 | 148.5 | 138.5 | 130.8271 | 121 | 141 | 86 | 148.9277 | 139 | 145
| Mean |  | 135 | 139.3333 | 143.3333 | 149.4167 | 141.1667 | 133.5994 | 120.6667 | 137.6667 | 86.2333 | 150.9246 | 139.1667 | 145.6667
| Median |  | 134.75 | 140.5 | 142.75 | 149.25 | 141.25 | 133.7286 | 120.5 | 138 | 85.9 | 149.4507 | 139 | 145.5
| Std Dev. |  | 1.5166 | 5.3635 | 2.0412 | 2.0595 | 3.2352 | 2.7491 | 4.9666 | 2.7325 | 2.3441 | 3.7502 | 0.40825 | 2.6583
| Rel.Std.Dev. |  | 1.123% | 3.849% | 1.424% | 1.378% | 2.292% | 2.058% | 4.116% | 1.985% | 2.718% | 2.485% | 0.2934% | 1.825%
| PDM3 |  | -4.513% | -1.448% | 1.382% | 5.685% | -0.1508% | -5.503% | -14.65% | -2.626% | -39.01% | 6.751% | -1.565% | 3.032%
| Z-Score (Absolute) |  | 0.58 | 0.01 | 0.56 | 1.38 | 0.26 | 0.77 | 2.53 | 0.22 | 7.22 | 1.39 | 0.01 | 0.87
| Analytical results for Se in OREAS 464 (Certified Value < 50 ppm) | [15 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 13 | < 50 | 16 | 27 | 40 | 10.1443 | 28 | 35 | < 0.5 | 5.7 | 30 |  |  |  | 1
| 1 | 2 | 13 | < 50 | 18 | 27 | 40 | 13.5092 | 28 | 38 | < 0.5 | 5.5 | 35 |  |  |  | 15
| 1 | 3 | 13 | < 50 | 18 | 27 | 39 | 11.986 | 29 | 38 | < 0.5 | 1.2 | 30 |  |  |  | 16
| 1 | 4 | 14 | < 50 | 17 | 26 | 39 | 12.8004 | 29 | 36 | < 0.5 | 6.4 | 30 |  |  |  | < 50
| 1 | 5 | 12 | < 50 | 17 | 27 | 41 | 9.7127 | 27 | 35 | < 0.5 | 4.5 | 35 |  |  |  | 107
| 1 | 6 | 13 | < 50 | 16 | 26 | 39 | 13.1495 | 29 | 38 | < 0.5 | 3.6 | 30
| Mean |  | 13 |  | 17 | 26.6667 | 39.6667 | 11.8837 | 28.3333 | 36.6667 |  | 4.4833 | 31.6667
| Median |  | 13 |  | 17 | 27 | 39.5 | 12.3932 | 28.5 | 37 |  | 5 | 30
| Std Dev. |  | 0.63246 |  | 0.89443 | 0.5164 | 0.8165 | 1.6021 | 0.8165 | 1.5055 |  | 1.8841 | 2.582
| Rel.Std.Dev. |  | 4.865% |  | 5.261% | 1.936% | 2.058% | 13.48% | 2.882% | 4.106% |  | 42.02% | 8.154%
PDM3
| Z-Score (Absolute) |  | 0.92 | NA | 0.65 | 0 | 0.88 | 1 | 0.11 | 0.67 | NA | 1.5 | 0.34
| Analytical results for Sm in OREAS 464 (Certified Value 1485 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1603.3667 | 1019.4 | 1470.4 | 748 | 1446.62 | 1425.2329 | 1449.8 | 1490 |  | 1
| 1 | 2 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1569.56 | 1097.5 | 1494.5 | 745 | 1471.25 | 1396.2436 | 1457.6 | 1480 |  | #N/A
| 1 | 3 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1493.8714 | 1089.6 | 1491.4 | 712 | 1496.2 | 1451.6107 | 1519.9 | 1520 |  | 16
| 1 | 4 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1534.734 | 1048.3 | 1474.2 | 715 | 1454.61 | 1403.0298 | 1444.5 | 1510 |  | 1484.8664
| 1 | 5 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1608.871 | 987.3 | 1476.6 | 741 | 1511.9 | 1408.1493 | 1461.7 | 1520 |  | 108
| 1 | 6 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1531.924 | 1081.6 | 1485 | 721 | 1430.31 | 1444.2079 | 1507.9 | 1520
| Mean | [6 empty] | 1557.0545 | 1053.95 | 1482.0167 | 730.3333 | 1468.4817 | 1421.4124 | 1473.5667 | 1506.6667
| Median | [6 empty] | 1552.147 | 1064.95 | 1480.8 | 731 | 1462.93 | 1416.6911 | 1459.65 | 1515
| Std Dev. | [6 empty] | 44.9622 | 43.8176 | 9.7792 | 16.1204 | 30.9469 | 22.7749 | 32.0345 | 17.5119
| Rel.Std.Dev. | [6 empty] | 2.888% | 4.157% | 0.6599% | 2.207% | 2.107% | 1.602% | 2.174% | 1.162%
| PDM3 | [6 empty] | 4.862% | -29.02% | -0.1919% | -50.81% | -1.103% | -4.273% | -0.761% | 1.468%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 1.36 | 6.6 | 0.17 | 11.72 | 0.04 | 0.78 | 0.04 | 0.56
| Analytical results for Sn in OREAS 464 (Certified Value 35.7 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 33.8 | 39.2 | 42 | 37.4 | 37.2 | 31.714 | 36.3 | 37 | 28.6 | 37.5 | 37 |  |  |  | 1
| 1 | 2 | 33.5 | 37.9 | 40.3 | 36.4 | 36.9 | 33.4239 | 36.4 | 33 | 29 | 34.8 | 37 |  |  |  | #N/A
| 1 | 3 | 34.4 | 35.1 | 41.5 | 36.5 | 36.7 | 31.4885 | 36.6 | 33 | 29.2 | 36 | 36 |  |  |  | 16
| 1 | 4 | 33.6 | 37.1 | 42.2 | 37.4 | 37.9 | 32.9356 | 36.3 | 35 | 28.5 | 36.4 | 36 |  |  |  | 35.6779
| 1 | 5 | 33.7 | 40.5 | 41.3 | 38.2 | 38.7 | 32.8752 | 35.8 | 32 | 29.9 | 35.7 | 37 |  |  |  | 109
| 1 | 6 | 34.3 | 39.1 | 39.9 | 36.9 | 37.9 | 32.3009 | 35.6 | 35 | 29.6 | 36.3 | 36
| Mean |  | 33.8833 | 38.15 | 41.2 | 37.1333 | 37.55 | 32.4564 | 36.1667 | 34.1667 | 29.1333 | 36.1167 | 36.5
| Median |  | 33.75 | 38.5 | 41.4 | 37.15 | 37.55 | 32.5881 | 36.3 | 34 | 29.1 | 36.15 | 36.5
| Std Dev. |  | 0.37639 | 1.8971 | 0.92087 | 0.67429 | 0.75299 | 0.75545 | 0.38297 | 1.8348 | 0.55015 | 0.88863 | 0.54772
| Rel.Std.Dev. |  | 1.111% | 4.973% | 2.235% | 1.816% | 2.005% | 2.328% | 1.059% | 5.37% | 1.888% | 2.46% | 1.501%
| PDM3 |  | -5.03% | 6.929% | 15.48% | 4.08% | 5.247% | -9.029% | 1.37% | -4.236% | -18.34% | 1.23% | 2.304%
| Z-Score (Absolute) |  | 0.78 | 0.67 | 1.71 | 0.33 | 0.47 | 1.26 | 0 | 0.68 | 2.39 | 0.02 | 0.11
| Analytical results for Sr in OREAS 464 (Certified Value 2043 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 2290 | 2200 | 1070 | 1150 | 951 | 2093 | 1890.7 | 2194.9 | 1820 | 2022.37 | 2048.6222 | 1949 | 2065 | 1820 | 1
| 1 | 2 | 2310 | 2160 | 1050 | 1130 | 960 | 2072.5164 | 1912.8 | 2215.3 | 1840 | 1987.56 | 2094.1348 | 1953 | 2070 | 1830 | #N/A
| 1 | 3 | 2300 | 2160 | 1060 | 1135 | 966 | 2021 | 1894.3 | 2228.5 | 1860 | 2075.14 | 2068.5169 | 1814 | 2070 | 1860 | 16
| 1 | 4 | 2320 | 2130 | 1060 | 1170 | 1010 | 2052.2607 | 1894.4 | 2188.7 | 1830 | 2019.1 | 2092.2163 | 1968 | 2065 | 1840 | 2042.8066
| 1 | 5 | 2280 | 2160 | 1050 | 1190 | 953 | 2032.9013 | 1863.8 | 2179.9 | 1850 | 2078.9 | 2098.7302 | 1948 | 2065 | 1870 | 110
| 1 | 6 | 2330 | 2200 | 1050 | 1175 | 980 | 2007.5389 | 1863.7 | 2188.5 | 1870 | 2050.96 | 2161.2653 | 1997 | 2070 | 1920
| Mean |  | 2305 | 2168.3333 | 1056.6667 | 1158.3333 | 970 | 2046.5362 | 1886.6167 | 2199.3 | 1845 | 2039.005 | 2093.9143 | 1938.1667 | 2067.5 | 1856.6667
| Median |  | 2305 | 2160 | 1055 | 1160 | 963 | 2042.581 | 1892.5 | 2191.8 | 1845 | 2036.665 | 2093.1755 | 1951 | 2067.5 | 1850
| Std Dev. |  | 18.7083 | 27.1416 | 8.165 | 23.8048 | 22.2081 | 32.3278 | 19.3268 | 18.6191 | 18.7083 | 35.6645 | 38.1016 | 63.5655 | 2.7386 | 36.1478
| Rel.Std.Dev. |  | 0.8116% | 1.252% | 0.7727% | 2.055% | 2.289% | 1.58% | 1.024% | 0.8466% | 1.014% | 1.749% | 1.82% | 3.28% | 0.1325% | 1.947%
| PDM3 |  | 12.83% | 6.145% | -48.27% | -43.3% | -52.52% | 0.1826% | -7.646% | 7.661% | -9.683% | -0.1861% | 2.502% | -5.122% | 1.209% | -9.112%
| Z-Score (Absolute) |  | 1.37 | 0.75 | 4.24 | 3.78 | 4.67 | 0.2 | 0.51 | 0.89 | 0.7 | 0.17 | 0.42 | 0.17 | 0.3 | 0.65
| Analytical results for Ta in OREAS 464 (Certified Value 20.9 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 24.1 | 22.9 | 22.5 | 22.9 | 19.1134 | 20.4 | 15.1 | 11.2 | 16.57 | 20.5 |  |  |  |  | 1
| 1 | 2 | 23.9 | 23 | 22.5 | 22 | 18.8427 | 21 | 10.4 | 11.1 | 13.41 | 20.9 |  |  |  |  | #N/A
| 1 | 3 | 24.3 | 22.6 | 22.2 | 22.4 | 18.2164 | 20.8 | 12.1 | 11.6 | 16.65 | 19.6 |  |  |  |  | 16
| 1 | 4 | 23.7 | 22.3 | 22.8 | 22.9 | 18.8374 | 21.6 | 12.6 | 11.7 | 16.91 | 20.1 |  |  |  |  | 20.9121
| 1 | 5 | 23.1 | 23.3 | 23 | 23.3 | 18.1333 | 21 | 8.5 | 11.5 | 14.76 | 21.3 |  |  |  |  | 111
| 1 | 6 | 23.8 | 22.2 | 22.1 | 23.1 | 19.3882 | 20.2 | 14.2 | 11.5 | 13.75 | 20.9
| Mean |  | 23.8167 | 22.7167 | 22.5167 | 22.7667 | 18.7552 | 20.8333 | 12.15 | 11.4333 | 15.3417 | 20.55
| Median |  | 23.85 | 22.75 | 22.5 | 22.9 | 18.8401 | 20.9 | 12.35 | 11.5 | 15.665 | 20.7
| Std Dev. |  | 0.41191 | 0.42622 | 0.34303 | 0.48028 | 0.49405 | 0.49666 | 2.4288 | 0.23381 | 1.5674 | 0.61887
| Rel.Std.Dev. |  | 1.729% | 1.876% | 1.523% | 2.11% | 2.634% | 2.384% | 19.99% | 2.045% | 10.22% | 3.012%
| PDM3 |  | 13.89% | 8.629% | 7.673% | 8.868% | -10.31% | -0.3767% | -41.9% | -45.33% | -26.64% | -1.732%
| Z-Score (Absolute) |  | 1.03 | 0.67 | 0.6 | 0.68 | 0.64 | 0.05 | 2.81 | 3.05 | 1.76 | 0.05
| Analytical results for Tb in OREAS 464 (Certified Value 55 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 54 | 56.3 | 50.9 | 53 | 59.1 | 54.3175 | 45.2 | 56.08 | 33.6 | 57.66 | 84.8365 | 53.05 | 53.9 |  | 1
| 1 | 2 | 54.3 | 55.4 | 51.1 | 52.1 | 57.7 | 54.13 | 46.5 | 56.78 | 33 | 57.61 | 81.8399 | 53.5 | 54.7 |  | #N/A
| 1 | 3 | 54.2 | 50 | 50.4 | 52.3 | 58.1 | 53.2218 | 46 | 56.43 | 34.6 | 59.31 | 80.7973 | 54.8 | 55.6 |  | 16
| 1 | 4 | 54.4 | 55.2 | 50.3 | 52.2 | 59.1 | 53.8758 | 45.8 | 56 | 32.5 | 57.76 | 82.2319 | 52.95 | 55.1 |  | 54.8891
| 1 | 5 | 54.8 | 55.2 | 50.5 | 53.4 | 61.2 | 54.249 | 44.6 | 55.74 | 32.9 | 57.81 | 82.9475 | 53.5 | 55.2 |  | 112
| 1 | 6 | 53.4 | 56.3 | 49.8 | 52.2 | 59.4 | 52.8394 | 47.8 | 57.43 | 33.5 | 58.17 | 81.7471 | 55.45 | 54.2
| Mean |  | 54.1833 | 54.7333 | 50.5 | 52.5333 | 59.1 | 53.7722 | 45.9833 | 56.41 | 33.35 | 58.0533 | 82.4001 | 53.875 | 54.7833
| Median |  | 54.25 | 55.3 | 50.45 | 52.25 | 59.1 | 54.0029 | 45.9 | 56.255 | 33.25 | 57.785 | 82.0359 | 53.5 | 54.9
| Std Dev. |  | 0.46655 | 2.3746 | 0.46043 | 0.53541 | 1.2215 | 0.60606 | 1.1071 | 0.61703 | 0.73417 | 0.64643 | 1.3837 | 1.0153 | 0.64317
| Rel.Std.Dev. |  | 0.8611% | 4.338% | 0.9118% | 1.019% | 2.067% | 1.127% | 2.408% | 1.094% | 2.201% | 1.114% | 1.679% | 1.884% | 1.174%
| PDM3 |  | -1.286% | -0.2837% | -7.996% | -4.292% | 7.672% | -2.035% | -16.22% | 2.771% | -39.24% | 5.765% | 50.12% | -1.847% | -0.1926%
| Z-Score (Absolute) |  | 0 | 0.45 | 1.12 | 0.5 | 1.49 | 0.12 | 2.48 | 0.67 | 6.31 | 1.17 | 8.54 | 0.09 | 0.18
| Analytical results for Te in OREAS 464 (Certified Value 0.36 ppm) | [15 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.43 | 0.31 | 0.49 | 0.34 | 0.34 | 0.27334 | < 1 | < 0.5 | 0.48 | 0.3 | 0.2 |  |  |  | 1
| 1 | 2 | 0.42 | 0.33 | 0.58 | 0.32 | 0.32 | 0.28112 | < 1 | < 0.5 | 0.46 | 0.4 | 0.2 |  |  |  | 16
| 1 | 3 | 0.4 | 0.34 | 0.58 | 0.4 | 0.33 | 0.24571 | < 1 | < 0.5 | 0.46 | 0.4 | 0.2 |  |  |  | 16
| 1 | 4 | 0.41 | 0.32 | 0.54 | 0.37 | 0.33 | 0.23765 | < 1 | < 0.5 | 0.46 | 0.3 | 0.2 |  |  |  | 0.3625
| 1 | 5 | 0.42 | 0.38 | 0.56 | 0.41 | 0.33 | 0.23405 | < 1 | < 0.5 | 0.49 | 0.3 | 0.2 |  |  |  | 113
| 1 | 6 | 0.43 | 0.36 | 0.47 | 0.41 | 0.34 | 0.273 | < 1 | < 0.5 | 0.44 | 0.4 | 0.2
| Mean |  | 0.41833 | 0.34 | 0.53667 | 0.375 | 0.33167 | 0.25748 |  |  | 0.465 | 0.35 | 0.2
| Median |  | 0.42 | 0.335 | 0.55 | 0.385 | 0.33 | 0.25936 |  |  | 0.46 | 0.35 | 0.2
| Std Dev. |  | 0.01169 | 0.026077 | 0.046762 | 0.038341 | 0.0075277 | 0.020648 |  |  | 0.017607 | 0.054772 | 3.0405e-17
| Rel.Std.Dev. |  | 2.795% | 7.67% | 8.713% | 10.22% | 2.27% | 8.019% |  |  | 3.786% | 15.65% | 1.52e-14%
| PDM3 |  | 15.4% | -6.206% | 48.05% | 3.449% | -8.505% | -28.97% |  |  | 28.28% | -3.447% | -44.83%
| Z-Score (Absolute) |  | 0.67 | 0.1 | 1.84 | 0.25 | 0.18 | 0.91 | NA | NA | 1.13 | 0 | 1.48
| Analytical results for Th in OREAS 464 (Certified Value 506 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 870 | 390 | 556 | 383 | 525 | 511.13 | 544.1 | 81.5 | 524.97 | 563.3469 | 517 |  |  |  | 1
| 1 | 2 | 840 | 405 | 534 | 386 | 508 | 508.5451 | 547 | 80.4 | 530.88 | 561.0541 | 524 |  |  |  | #N/A
| 1 | 3 | 870 | 410 | 542 | 405 | 517 | 498.5196 | 564.4 | 85.4 | 549.57 | 562.7466 | 537 |  |  |  | 16
| 1 | 4 | 880 | 390 | 540 | 406 | 529 | 508.1129 | 551.6 | 83.6 | 522.57 | 563.8497 | 536 |  |  |  | 506.3839
| 1 | 5 | 850 | 419 | 563 | 413 | 543 | 512.1607 | 536.7 | 82.9 | 536.57 | 564.5312 | 531 |  |  |  | 114
| 1 | 6 | 860 | 405 | 546 | 402 | 528 | 498.5466 | 552.5 | 87.3 | 550.8 | 563.5296 | 527
| Mean |  | 861.6667 | 403.1667 | 546.8333 | 399.1667 | 525 | 506.1692 | 549.3833 | 83.5167 | 535.8933 | 563.1763 | 528.6667
| Median |  | 865 | 405 | 544 | 403.5 | 526.5 | 508.329 | 549.3 | 83.25 | 533.725 | 563.4383 | 529
| Std Dev. |  | 14.7196 | 11.4091 | 10.7781 | 11.9569 | 11.8491 | 6.1082 | 9.323 | 2.5309 | 12.097 | 1.1943 | 7.607
| Rel.Std.Dev. |  | 1.708% | 2.83% | 1.971% | 2.995% | 2.257% | 1.207% | 1.697% | 3.03% | 2.257% | 0.2121% | 1.439%
| PDM3 |  | 70.16% | -20.38% | 7.988% | -21.17% | 3.676% | -0.04242% | 8.491% | -83.51% | 5.827% | 11.22% | 4.4%
| Z-Score (Absolute) |  | 9.98 | 3.76 | 0.54 | 3.88 | 0.11 | 0.67 | 0.62 | 13.34 | 0.22 | 1.03 | 0
| Analytical results for Ti in OREAS 464 (Certified Value < 1 wt.%) | [15 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.899 | 0.89 | 0.462 | 0.412 | 0.383 | 0.71148 | 0.67 | 0.7085 | 0.656 | 0.891 | 0.55 | 1.09 |  |  | 1
| 1 | 2 | 0.863 | 0.87 | 0.448 | 0.402 | 0.367 | 0.71117 | 0.68 | 0.6161 | 0.674 | 0.8842 | 0.55 | 1.1 |  |  | #N/A
| 1 | 3 | 0.895 | 0.9 | 0.453 | 0.397 | 0.392 | 0.7496 | 0.68 | 0.5938 | 0.667 | 0.9083 | 0.52 | 1.14 |  |  | 16
| 1 | 4 | 0.871 | 0.89 | 0.426 | 0.415 | 0.401 | 0.74343 | 0.69 | 0.6138 | 0.681 | 0.9009 | 0.52 | 1.2 |  |  | < 1
| 1 | 5 | 0.837 | 0.93 | 0.446 | 0.42 | 0.369 | 0.6888 | 0.68 | 0.5722 | 0.685 | 0.856 | 0.53 | 1.17 |  |  | 115
| 1 | 6 | 0.866 | 0.92 | 0.428 | 0.422 | 0.401 | 0.74753 | 0.66 | 0.7105 | 0.679 | 0.8493 | 0.54 | 1.14
| Mean |  | 0.87183 | 0.9 | 0.44383 | 0.41133 | 0.3855 | 0.72533 | 0.67667 | 0.63582 | 0.67367 | 0.88162 | 0.535 | 1.14
| Median |  | 0.8685 | 0.895 | 0.447 | 0.4135 | 0.3875 | 0.72746 | 0.68 | 0.61495 | 0.6765 | 0.8876 | 0.535 | 1.14
| Std Dev. |  | 0.022807 | 0.021909 | 0.014176 | 0.0099532 | 0.015123 | 0.025045 | 0.010328 | 0.059239 | 0.010652 | 0.023996 | 0.013784 | 0.041473
| Rel.Std.Dev. |  | 2.616% | 2.434% | 3.194% | 2.42% | 3.923% | 3.453% | 1.526% | 9.317% | 1.581% | 2.722% | 2.576% | 3.638%
PDM3
| Z-Score (Absolute) |  | 0.66 | 0.75 | 0.77 | 0.88 | 0.97 | 0.17 | 0.01 | 0.13 | 0.01 | 0.69 | 0.47 | 1.56
| Analytical results for Tl in OREAS 464 (Certified Value 0.15 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.15 | 0.09 | 0.15 | 0.17 | 0.15 | 0.1429 | 0.14 | < 0.2 | 0.14 | 0.13 | 0.16333 | 0.15 | 0.1 |  | 1
| 1 | 2 | 0.15 | 0.1 | 0.16 | 0.15 | 0.14 | 0.1357 | 0.14 | 0.2 | 0.14 | 0.12 | 0.21759 | 0.15 | 0.1 |  | 17
| 1 | 3 | 0.15 | 0.05 | 0.15 | 0.17 | 0.14 | 0.15358 | 0.15 | < 0.2 | 0.15 | 0.11 | 0.21743 | 0.15 | 0.1 |  | 16
| 1 | 4 | 0.15 | 0.09 | 0.14 | 0.16 | 0.15 | 0.12752 | 0.15 | < 0.2 | 0.15 | 0.13 | 0.2291 | 0.15 | 0.1 |  | 0.14535
| 1 | 5 | 0.16 | 0.09 | 0.14 | 0.17 | 0.15 | 0.13365 | 0.15 | < 0.2 | 0.14 | 0.12 | 0.19834 | 0.15 | 0.1 |  | 116
| 1 | 6 | 0.16 | 0.08 | 0.13 | 0.15 | 0.15 | 0.15551 | 0.15 | < 0.2 | 0.14 | 0.11 | 0.21758 | 0.15 | 0.1
| Mean |  | 0.15333 | 0.083333 | 0.145 | 0.16167 | 0.14667 | 0.14147 | 0.14667 | 0.2 | 0.14333 | 0.12 | 0.20723 | 0.15 | 0.1
| Median |  | 0.15 | 0.09 | 0.145 | 0.165 | 0.15 | 0.1393 | 0.15 | 0.2 | 0.14 | 0.12 | 0.21751 | 0.15 | 0.1
| Std Dev. |  | 0.005164 | 0.017512 | 0.010488 | 0.0098319 | 0.005164 | 0.011269 | 0.005164 |  | 0.005164 | 0.0089443 | 0.023676 | 0 | 1.5202e-17
| Rel.Std.Dev. |  | 3.368% | 21.01% | 7.233% | 6.082% | 3.521% | 7.965% | 3.521% |  | 3.603% | 7.454% | 11.43% | 0% | 1.52e-14%
| PDM3 |  | 5.493% | -42.67% | -0.2402% | 11.23% | 0.9065% | -2.665% | 0.9065% | 37.6% | -1.387% | -17.44% | 42.57% | 3.2% | -31.2%
| Z-Score (Absolute) |  | 0.67 | 5.73 | 0.17 | 1.52 | 0 | 0.53 | 0 | 5.39 | 0.34 | 2.7 | 7.01 | 0.34 | 4.72
| Analytical results for Tm in OREAS 464 (Certified Value 3.3 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.84 | 3.63 | 3.27 | 3.3 | 3.52 | 3.087 | 2.6545 | 3.7 | 1.87 | 2.97 | 4.3939 | 3.4 | 3.02 |  | 1
| 1 | 2 | 3.75 | 3.61 | 3.25 | 3.29 | 3.46 | 3.2078 | 2.704 | 3.7 | 1.85 | 2.96 | 4.2954 | 3.35 | 3.04 |  | #N/A
| 1 | 3 | 3.77 | 3.21 | 3.21 | 3.26 | 3.45 | 3.0746 | 2.7742 | 3.8 | 1.97 | 2.92 | 4.1532 | 3.4 | 3.08 |  | 16
| 1 | 4 | 3.79 | 3.44 | 3.2 | 3.25 | 3.54 | 3.0566 | 2.6781 | 3.5 | 1.78 | 2.92 | 4.2428 | 3.35 | 3.16 |  | 3.2956
| 1 | 5 | 3.76 | 3.27 | 3.3 | 3.39 | 3.66 | 3.261 | 2.6198 | 3.6 | 1.82 | 3.09 | 4.2235 | 3.4 | 3.16 |  | 117
| 1 | 6 | 3.83 | 3.61 | 3.16 | 3.26 | 3.52 | 3.13 | 2.6725 | 3.7 | 1.75 | 2.88 | 4.3145 | 3.45 | 3.16
| Mean |  | 3.79 | 3.4617 | 3.2317 | 3.2917 | 3.525 | 3.1362 | 2.6839 | 3.6667 | 1.84 | 2.9567 | 4.2706 | 3.3917 | 3.1033
| Median |  | 3.78 | 3.525 | 3.23 | 3.275 | 3.52 | 3.1085 | 2.6753 | 3.7 | 1.835 | 2.94 | 4.2691 | 3.4 | 3.12
| Std Dev. |  | 0.037417 | 0.18595 | 0.051153 | 0.051929 | 0.075299 | 0.0816 | 0.052328 | 0.10328 | 0.07746 | 0.072847 | 0.083074 | 0.037639 | 0.065013
| Rel.Std.Dev. |  | 0.9872% | 5.372% | 1.583% | 1.578% | 2.136% | 2.602% | 1.95% | 2.817% | 4.21% | 2.464% | 1.945% | 1.11% | 2.095%
| PDM3 |  | 15% | 5.038% | -1.941% | -0.1205% | 6.96% | -4.839% | -18.56% | 11.26% | -44.17% | -10.29% | 29.58% | 2.914% | -5.835%
| Z-Score (Absolute) |  | 1.38 | 0.51 | 0.11 | 0 | 0.67 | 0.36 | 1.57 | 1.14 | 3.82 | 0.84 | 2.66 | 0.32 | 0.45
| Analytical results for U in OREAS 464 (Certified Value 17.6 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 17.2 | 17.9 | 17.8 | 18.4 | 19.6 | 16.6327 | 18.13 | 18.1 | 17.9 | 17.11 | 22.176 | 17.3 | 16.2 |  | 1
| 1 | 2 | 16.8 | 18.5 | 17.2 | 18.1 | 19.3 | 16.4923 | 18.13 | 18.2 | 18.1 | 17.4 | 21.2595 | 17.2 | 15.2 |  | #N/A
| 1 | 3 | 17.1 | 15.6 | 17.6 | 18 | 19.2 | 16.1846 | 18.11 | 18.1 | 18.6 | 17.44 | 21.6651 | 18 | 16.5 |  | 16
| 1 | 4 | 17.4 | 17.1 | 17.1 | 18 | 19.6 | 16.2618 | 18.76 | 18 | 17.9 | 17.15 | 21.6444 | 17.4 | 16.4 |  | 17.6356
| 1 | 5 | 17.1 | 18.1 | 18.3 | 18.7 | 20 | 16.5144 | 18.33 | 17.8 | 18.5 | 17.94 | 21.6024 | 17.1 | 15.7 |  | 118
| 1 | 6 | 17.6 | 17.6 | 17.4 | 18.1 | 19.4 | 16.4574 | 17.52 | 18.1 | 18.1 | 17.31 | 22.0646 | 17.5 | 15.6
| Mean |  | 17.2 | 17.4667 | 17.5667 | 18.2167 | 19.5167 | 16.4239 | 18.1633 | 18.05 | 18.1833 | 17.3917 | 21.7353 | 17.4167 | 15.9333
| Median |  | 17.15 | 17.75 | 17.5 | 18.1 | 19.5 | 16.4748 | 18.13 | 18.1 | 18.1 | 17.355 | 21.6548 | 17.35 | 15.95
| Std Dev. |  | 0.27568 | 1.0289 | 0.44121 | 0.27869 | 0.28577 | 0.16801 | 0.40068 | 0.13784 | 0.29944 | 0.29916 | 0.33461 | 0.31885 | 0.51251
| Rel.Std.Dev. |  | 1.603% | 5.891% | 2.512% | 1.53% | 1.464% | 1.023% | 2.206% | 0.7637% | 1.647% | 1.72% | 1.539% | 1.831% | 3.217%
| PDM3 |  | -2.47% | -0.9577% | -0.3906% | 3.295% | 10.67% | -6.871% | 2.993% | 2.35% | 3.106% | -1.383% | 23.25% | -1.241% | -9.652%
| Z-Score (Absolute) |  | 1.02 | 0 | 0.44 | 0.45 | 2.67 | 2.26 | 0.52 | 0.33 | 0.55 | 0.71 | 6.2 | 0.67 | 3.04
| Analytical results for V in OREAS 464 (Certified Value 207 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 199 | 187 | 111 | 117 | 106 | 204.9886 | 217.6 | 229 | 210 | 253 | 205.715 | 142 | 220 |  | 1
| 1 | 2 | 201 | 191 | 110 | 117 | 108 | 205.7898 | 223 | 220 | 209 | 258 | 207.297 | 149 | 220 |  | #N/A
| 1 | 3 | 198 | 190 | 110 | 117 | 107 | 201.1238 | 223.2 | 232 | 213 | 248 | 207.06 | 139 | 225 |  | 16
| 1 | 4 | 203 | 194 | 110 | 118 | 112 | 204.4475 | 219.5 | 228 | 202 | 243 | 208.287 | 147 | 230 |  | 207.0274
| 1 | 5 | 195 | 190 | 109 | 122 | 107 | 202.3107 | 217.5 | 226 | 205 | 239 | 207.559 | 145 | 225 |  | 119
| 1 | 6 | 204 | 206 | 111 | 119 | 109 | 199.965 | 215 | 225 | 206 | 242 | 207.898 | 151 | 220
| Mean |  | 200 | 193 | 110.1667 | 118.3333 | 108.1667 | 203.1042 | 219.3 | 226.6667 | 207.5 | 247.1667 | 207.3027 | 145.5 | 223.3333
| Median |  | 200 | 190.5 | 110 | 117.5 | 107.5 | 203.3791 | 218.55 | 227 | 207.5 | 245.5 | 207.428 | 146 | 222.5
| Std Dev. |  | 3.3466 | 6.7528 | 0.75277 | 1.9664 | 2.137 | 2.3227 | 3.2729 | 4.0825 | 3.937 | 7.2503 | 0.89091 | 4.4609 | 4.0825
| Rel.Std.Dev. |  | 1.673% | 3.499% | 0.6833% | 1.662% | 1.976% | 1.144% | 1.492% | 1.801% | 1.897% | 2.933% | 0.4298% | 3.066% | 1.828%
| PDM3 |  | -3.394% | -6.776% | -46.79% | -42.84% | -47.75% | -1.895% | 5.928% | 9.486% | 0.2283% | 19.39% | 0.133% | -29.72% | 7.876%
| Z-Score (Absolute) |  | 0.1 | 0.42 | 3.1 | 2.85 | 3.16 | 0 | 0.54 | 0.79 | 0.15 | 1.47 | 0.14 | 1.92 | 0.67
| Analytical results for W in OREAS 464 (Certified Value 2.37 ppm) | [15 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.9 | 2.5 | 3.2 | 2.7 | 2.7 | 2.2315 | 2 | 2 | 2.1 | 1.8 | 2 |  |  |  | 1
| 1 | 2 | 2.8 | 2.5 | 3.2 | 2.7 | 2.6 | 2.2597 | 2 | 2 | 1.9 | 1.9 | 2 |  |  |  | #N/A
| 1 | 3 | 2.8 | 2.3 | 3.1 | 2.5 | 2.7 | 2.0179 | 2 | 2 | 2.1 | 2.1 | 2.5 |  |  |  | 16
| 1 | 4 | 2.8 | 2.6 | 3 | 2.6 | 2.7 | 1.9838 | 2 | 2 | 2 | 1.9 | 2.5 |  |  |  | 2.3747
| 1 | 5 | 2.8 | 2.2 | 3.2 | 2.8 | 2.7 | 2.0233 | 2 | 2 | 2.1 | 1.8 | 2.5 |  |  |  | 120
| 1 | 6 | 2.8 | 2.5 | 3 | 2.7 | 2.7 | 2.3131 | 2 | 2 | 2.1 | 1.8 | 2.5
| Mean |  | 2.8167 | 2.4333 | 3.1167 | 2.6667 | 2.6833 | 2.1382 | 2 | 2 | 2.05 | 1.8833 | 2.3333
| Median |  | 2.8 | 2.5 | 3.15 | 2.7 | 2.7 | 2.1274 | 2 | 2 | 2.1 | 1.85 | 2.5
| Std Dev. |  | 0.040825 | 0.15055 | 0.098319 | 0.10328 | 0.040825 | 0.14534 | 0 | 0 | 0.083666 | 0.1169 | 0.2582
| Rel.Std.Dev. |  | 1.449% | 6.187% | 3.155% | 3.873% | 1.521% | 6.797% | 0% | 0% | 4.081% | 6.207% | 11.07%
| PDM3 |  | 18.61% | 2.47% | 31.25% | 12.3% | 13% | -9.959% | -15.78% | -15.78% | -13.67% | -20.69% | -1.741%
| Z-Score (Absolute) |  | 0.98 | 0.2 | 1.58 | 0.67 | 0.71 | 0.39 | 0.67 | 0.67 | 0.57 | 0.91 | 0
| Analytical results for Y in OREAS 464 (Certified Value 449 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 421 | 403 | 470 | 222 | 467 | 434.2993 | 262 | 470.1 | 157 | 449.04 | 516.4307 | 418.7 | 425 |  | 1
| 1 | 2 | 421 | 395 | 477 | 220 | 464 | 451.2891 | 261 | 464.2 | 152 | 449.43 | 516.4215 | 419.2 | 436 |  | #N/A
| 1 | 3 | 420 | 380 | 462 | 221 | 456 | 449.883 | 282 | 471.7 | 166 | 463.29 | 513.8517 | 435.5 | 442 |  | 16
| 1 | 4 | 420 | 385 | 467 | 226 | 471 | 430.9344 | 267 | 461.9 | 157 | 453.05 | 517.5494 | 416.8 | 435 |  | 448.8679
| 1 | 5 | 421 | 391 | 474 | 230 | 489 | 427.3796 | 252 | 467.5 | 155 | 468.73 | 518.6566 | 424.7 | 439 |  | 121
| 1 | 6 | 421 | 412 | 454 | 225 | 469 | 429.2373 | 282 | 463.6 | 166 | 455.7 | 516.0025 | 432 | 438
| Mean |  | 420.6667 | 394.3333 | 467.3333 | 224 | 469.3333 | 437.1705 | 267.6667 | 466.5 | 158.8333 | 456.54 | 516.4854 | 424.4833 | 435.8333
| Median |  | 421 | 393 | 468.5 | 223.5 | 468 | 432.6169 | 264.5 | 465.85 | 157 | 454.375 | 516.4261 | 421.95 | 437
| Std Dev. |  | 0.5164 | 11.7587 | 8.3825 | 3.7417 | 10.9666 | 10.6477 | 12.1106 | 3.8951 | 5.8452 | 7.9227 | 1.6139 | 7.7241 | 5.8452
| Rel.Std.Dev. |  | 0.1228% | 2.982% | 1.794% | 1.67% | 2.337% | 2.436% | 4.525% | 0.835% | 3.68% | 1.735% | 0.3125% | 1.82% | 1.341%
| PDM3 |  | -6.283% | -12.15% | 4.114% | -50.1% | 4.559% | -2.606% | -40.37% | 3.928% | -64.61% | 1.709% | 15.06% | -5.432% | -2.904%
| Z-Score (Absolute) |  | 0.32 | 0.89 | 0.67 | 4.53 | 0.72 | 0.03 | 3.6 | 0.66 | 5.93 | 0.44 | 1.73 | 0.24 | 0
| Analytical results for Yb in OREAS 464 (Certified Value 14.3 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 14.95 | 14.6 | 13.55 | 13.8 | 14.1 | 13.5134 | 10.6 | 15.2 | 10.2 | 13.21 | 19.5926 | 15.5 | 14.7 |  | 1
| 1 | 2 | 14.45 | 14.4 | 13.75 | 13.5 | 14 | 13.18 | 11.3377 | 15.7 | 10.1 | 13.21 | 18.3324 | 15.65 | 14.8 |  | #N/A
| 1 | 3 | 14.65 | 12.75 | 13.3 | 13.5 | 13.85 | 12.638 | 11.3059 | 15.2 | 10.8 | 13.49 | 18.259 | 16.2 | 15 |  | 16
| 1 | 4 | 14.4 | 14.1 | 13.15 | 13.45 | 14.3 | 12.9755 | 11.0977 | 15.7 | 10.7 | 13.11 | 18.4202 | 15.65 | 15.2 |  | 14.277
| 1 | 5 | 14.35 | 14.2 | 13.85 | 13.85 | 14.7 | 13.3024 | 10.5895 | 15.5 | 10 | 14 | 18.5084 | 16.05 | 14.8 |  | 122
| 1 | 6 | 14.8 | 14.55 | 13.65 | 13.75 | 14.1 | 12.8498 | 10.9526 | 15.7 | 10.7 | 13.12 | 19.6124 | 16 | 14.9
| Mean |  | 14.6 | 14.1 | 13.5417 | 13.6417 | 14.175 | 13.0765 | 10.9806 | 15.5 | 10.4167 | 13.3567 | 18.7875 | 15.8417 | 14.9
| Median |  | 14.55 | 14.3 | 13.6 | 13.625 | 14.1 | 13.0777 | 11.0252 | 15.6 | 10.45 | 13.21 | 18.4643 | 15.825 | 14.85
| Std Dev. |  | 0.24083 | 0.6892 | 0.2691 | 0.17725 | 0.29623 | 0.31829 | 0.33037 | 0.24495 | 0.35449 | 0.34396 | 0.63685 | 0.27824 | 0.17889
| Rel.Std.Dev. |  | 1.65% | 4.888% | 1.987% | 1.299% | 2.09% | 2.434% | 3.009% | 1.58% | 3.403% | 2.575% | 3.39% | 1.756% | 1.201%
| PDM3 |  | 2.263% | -1.239% | -5.15% | -4.45% | -0.7141% | -8.408% | -23.09% | 8.567% | -27.04% | -6.446% | 31.59% | 10.96% | 4.364%
| Z-Score (Absolute) |  | 0.42 | 0.24 | 0.42 | 0.34 | 0 | 0.8 | 2.47 | 1.15 | 2.93 | 0.67 | 3.78 | 1.42 | 0.66
| Analytical results for Zn in OREAS 464 (Certified Value 1128 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 1080 | 1220 | 516 | 528 | 467 | 1160.7229 | 1087.2889 | 1173 | 1110 | 1094 | 1204.47 | 1010 | 1115 | 1090 | 1
| 1 | 2 | 1100 | 1210 | 504 | 520 | 469 | 1171.2183 | 1135.4791 | 1155 | 1110 | 1080 | 1197.1692 | 1108 | 1110 | 1080 | #N/A
| 1 | 3 | 1080 | 1230 | 514 | 526 | 474 | 1180.2368 | 1109.2882 | 1170 | 1100 | 1062 | 1192.7196 | 1058 | 1115 | 1090 | 16
| 1 | 4 | 1090 | 1280 | 506 | 536 | 472 | 1175.7576 | 1107.2715 | 1174 | 1110 | 1080 | 1204.686 | 1068 | 1110 | 1090 | 1128.0782
| 1 | 5 | 1050 | 1200 | 504 | 548 | 471 | 1164.2548 | 1097.9691 | 1152 | 1100 | 1069 | 1194.8688 | 1022 | 1115 | 1100 | 123
| 1 | 6 | 1080 | 1280 | 507 | 540 | 488 | 1176.2111 | 1079.9697 | 1214 | 1110 | 1101 | 1278.0288 | 1099 | 1110 | 1110
| Mean |  | 1080 | 1236.6667 | 508.5 | 533 | 473.5 | 1171.4002 | 1102.8778 | 1173 | 1106.6667 | 1081 | 1211.9904 | 1060.8333 | 1112.5 | 1093.3333
| Median |  | 1080 | 1225 | 506.5 | 532 | 471.5 | 1173.488 | 1102.6203 | 1171.5 | 1110 | 1080 | 1200.8196 | 1063 | 1112.5 | 1090
| Std Dev. |  | 16.7332 | 35.0238 | 5.2058 | 10.2567 | 7.5033 | 7.554 | 19.5711 | 22.163 | 5.164 | 14.6697 | 32.7266 | 39.5799 | 2.7386 | 10.328
| Rel.Std.Dev. |  | 1.549% | 2.832% | 1.024% | 1.924% | 1.585% | 0.6449% | 1.775% | 1.889% | 0.4666% | 1.357% | 2.7% | 3.731% | 0.2462% | 0.9446%
| PDM3 |  | -4.262% | 9.626% | -54.92% | -52.75% | -58.03% | 3.84% | -2.234% | 3.982% | -1.898% | -4.173% | 7.439% | -5.961% | -1.381% | -3.08%
| Z-Score (Absolute) |  | 0.23 | 1.8 | 7.65 | 7.33 | 8.14 | 0.95 | 0.06 | 0.87 | 0.11 | 0.22 | 1.31 | 0.48 | 0.19 | 0.06
| Analytical results for Zr in OREAS 464 (Certified Value 210 ppm) | [15 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 243 | 221 | 250 | 244 | 237 | 187.7672 | 216 | 104 | 186.7 | 122 | 110 | 407 |  |  | 1
| 1 | 2 | 238 | 222 | 244 | 240 | 236 | 189.4418 | 202 | 108 | 170.8 | 118 | 95 | 412 |  |  | #N/A
| 1 | 3 | 241 | 210 | 234 | 240 | 231 | 187.5 | 207 | 103 | 186.8 | 119 | 95 | 406 |  |  | 16
| 1 | 4 | 240 | 222 | 234 | 246 | 240 | 188.4863 | 204 | 106 | 192.8 | 115 | 160 | 406 |  |  | 209.7096
| 1 | 5 | 234 | 234 | 243 | 251 | 246 | 186.582 | 196 | 107 | 188.8 | 106 | 100 | 409 |  |  | 124
| 1 | 6 | 238 | 233 | 231 | 242 | 238 | 189.3831 | 218 | 109 | 181.7 | 117 | 165 | 419
| Mean |  | 239 | 223.6667 | 239.3333 | 243.8333 | 238 | 188.1934 | 207.1667 | 106.1667 | 184.6 | 116.1667 | 120.8333 | 409.8333
| Median |  | 239 | 222 | 238.5 | 243 | 237.5 | 188.1267 | 205.5 | 106.5 | 186.75 | 117.5 | 105 | 408
| Std Dev. |  | 3.0984 | 8.8694 | 7.4207 | 4.2151 | 4.9396 | 1.1242 | 8.4479 | 2.3166 | 7.6548 | 5.4924 | 32.7745 | 5.0365
| Rel.Std.Dev. |  | 1.296% | 3.965% | 3.101% | 1.729% | 2.075% | 0.5973% | 4.078% | 2.182% | 4.147% | 4.728% | 27.12% | 1.229%
| PDM3 |  | 13.97% | 6.655% | 14.13% | 16.27% | 13.49% | -10.26% | -1.213% | -49.37% | -11.97% | -44.61% | -42.38% | 95.43%
| Z-Score (Absolute) |  | 0.58 | 0.2 | 0.58 | 0.69 | 0.55 | 0.66 | 0.2 | 2.67 | 0.68 | 2.37 | 2.31 | 4.74
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.