## Abbreviations
Table 1. Abbreviations used for OREAS 610
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with AAS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
ABL*MS | Laser Ablation with Mass Spectrometry: ICP-MS finish
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/AAS | ore grade aqua regia digest with either AAS or ICP-OES finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
FA*AAS | fire assay with AAS finish
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
LOI*TGA | loss on ignition with Thermal Gravimetric Analyser finish
Raw*XRPA | raw sample (solid) packed into single use jar by X-Ray Photon Assay
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 610
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Chrysos Corporation Limited, Kalgoorlie, WA, Australia
Chrysos Corporation Limited, Perth, WA, Australia
Inspectorate (BV), Lima, Peru
Inspectorate America Corporation (BV), Sparks, Nevada, USA
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
Skyline Assayers & Laboratories, Tucson, Arizona, USA

## Performance Gates
Table 1. Pooled-Lab Performance Gates for OREAS 610
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 9.8341 | 0.25443 | 9.3252 | 10.343 | 9.0708 | 10.5974 | 2.587% | 5.175% | 7.762% | 9.3424 | 10.3258
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 9.8119 | 0.34094 | 9.13 | 10.4938 | 8.7891 | 10.8347 | 3.475% | 6.95% | 10.42% | 9.3213 | 10.3025
Infrared Combustion
S, wt.% | 4.1172 | 0.15678 | 3.8036 | 4.4307 | 3.6468 | 4.5875 | 3.808% | 7.616% | 11.42% | 3.9113 | 4.323
4-Acid Digestion
Ag, ppm | 49.3686 | 1.7899 | 45.7888 | 52.9484 | 43.9989 | 54.7383 | 3.626% | 7.251% | 10.88% | 46.9002 | 51.837
Al, wt.% | 5.9878 | 0.26049 | 5.4669 | 6.5088 | 5.2064 | 6.7693 | 4.35% | 8.701% | 13.05% | 5.6884 | 6.2872
As, ppm | 2834.5718 | 185.1685 | 2464.2348 | 3204.9087 | 2279.0664 | 3390.0772 | 6.533% | 13.07% | 19.6% | 2692.8432 | 2976.3004
Be, ppm | 1.4751 | 0.10935 | 1.2564 | 1.6938 | 1.147 | 1.8031 | 7.413% | 14.83% | 22.24% | 1.4013 | 1.5488
Bi, ppm | 223.5764 | 9.6809 | 204.2146 | 242.9381 | 194.5337 | 252.619 | 4.33% | 8.66% | 12.99% | 212.3976 | 234.7552
Ca, wt.% | 0.24074 | 0.014085 | 0.21257 | 0.26891 | 0.19849 | 0.28299 | 5.85% | 11.7% | 17.55% | 0.2287 | 0.25278
Cd, ppm | 12.0867 | 0.70848 | 10.6698 | 13.5037 | 9.9613 | 14.2122 | 5.862% | 11.72% | 17.58% | 11.4824 | 12.6911
Ce, ppm | 46.6884 | 4.4632 | 37.7621 | 55.6148 | 33.2989 | 60.078 | 9.559% | 19.12% | 28.68% | 44.354 | 49.0229
Co, ppm | 7.7238 | 0.30497 | 7.1138 | 8.3337 | 6.8089 | 8.6387 | 3.948% | 7.897% | 11.85% | 7.3376 | 8.11
Cr, ppm | 39.1331 | 5.0643 | 29.0046 | 49.2616 | 23.9403 | 54.3258 | 12.94% | 25.88% | 38.82% | 37.1764 | 41.0897
Cs, ppm | 2.1551 | 0.14615 | 1.8628 | 2.4474 | 1.7166 | 2.5935 | 6.782% | 13.56% | 20.34% | 2.0473 | 2.2629
Cu, wt.% | 0.97124 | 0.022851 | 0.92554 | 1.0169 | 0.90269 | 1.0398 | 2.353% | 4.705% | 7.058% | 0.92268 | 1.0198
Dy, ppm | 1.525 | 0.08893 | 1.3471 | 1.7029 | 1.2582 | 1.7918 | 5.832% | 11.66% | 17.49% | 1.4487 | 1.6013
Er, ppm | 0.59367 | 0.040994 | 0.51168 | 0.67566 | 0.47068 | 0.71665 | 6.905% | 13.81% | 20.72% | 0.56398 | 0.62335
Eu, ppm | 0.77133 | 0.050769 | 0.6698 | 0.87287 | 0.61903 | 0.92364 | 6.582% | 13.16% | 19.75% | 0.73277 | 0.8099
Fe, wt.% | 2.3658 | 0.076692 | 2.2125 | 2.5192 | 2.1358 | 2.5959 | 3.242% | 6.483% | 9.725% | 2.2476 | 2.4841
Ga, ppm | 24.419 | 1.0235 | 22.372 | 26.4661 | 21.3484 | 27.4897 | 4.192% | 8.383% | 12.57% | 23.1981 | 25.64
Gd, ppm | 2.856 | 0.25636 | 2.3433 | 3.3687 | 2.0869 | 3.6251 | 8.976% | 17.95% | 26.93% | 2.7132 | 2.9988
Hf, ppm | 2.0068 | 0.12755 | 1.7517 | 2.2619 | 1.6242 | 2.3895 | 6.356% | 12.71% | 19.07% | 1.9065 | 2.1071
Ho, ppm | 0.217 | 0.021359 | 0.17428 | 0.25972 | 0.15292 | 0.28108 | 9.843% | 19.69% | 29.53% | 0.20615 | 0.22785
In, ppm | 3.8952 | 0.24442 | 3.4064 | 4.384 | 3.1619 | 4.6284 | 6.275% | 12.55% | 18.82% | 3.7004 | 4.0899
K, wt.% | 1.9786 | 0.067443 | 1.8437 | 2.1135 | 1.7763 | 2.181 | 3.409% | 6.817% | 10.23% | 1.8797 | 2.0776
La, ppm | 20.3844 | 4.4167 | 11.551 | 29.2177 | 7.1344 | 33.6344 | 21.67% | 43.33% | 65% | 19.3652 | 21.4036
Li, ppm | 28.4423 | 1.3454 | 25.7515 | 31.1331 | 24.406 | 32.4785 | 4.73% | 9.461% | 14.19% | 27.0202 | 29.8644
Mg, ppm | 1557.4306 | 90.5031 | 1376.4243 | 1738.4369 | 1285.9212 | 1828.94 | 5.811% | 11.62% | 17.43% | 1479.5591 | 1635.3021
Mn, ppm | 77.7201 | 3.5584 | 70.6034 | 84.8368 | 67.045 | 88.3952 | 4.578% | 9.157% | 13.74% | 73.8341 | 81.6061
Mo, ppm | 4.8165 | 0.33191 | 4.1527 | 5.4803 | 3.8208 | 5.8122 | 6.891% | 13.78% | 20.67% | 4.5757 | 5.0573
Na, wt.% | 0.82035 | 0.026705 | 0.76694 | 0.87376 | 0.74024 | 0.90046 | 3.255% | 6.511% | 9.766% | 0.77933 | 0.86137
Nb, ppm | 8.2736 | 0.42983 | 7.4139 | 9.1332 | 6.9841 | 9.5631 | 5.195% | 10.39% | 15.59% | 7.8599 | 8.6872
Nd, ppm | 18.881 | 1.5596 | 15.7618 | 22.0002 | 14.2022 | 23.5598 | 8.26% | 16.52% | 24.78% | 17.9369 | 19.8251
Ni, ppm | 24.6174 | 1.1675 | 22.2824 | 26.9524 | 21.1149 | 28.1199 | 4.743% | 9.485% | 14.23% | 23.3865 | 25.8483
P, ppm | 554.2234 | 36.413 | 481.3975 | 627.0493 | 444.9846 | 663.4623 | 6.57% | 13.14% | 19.71% | 526.5123 | 581.9346
Pb, ppm | 661.6457 | 42.4928 | 576.66 | 746.6314 | 534.1672 | 789.1242 | 6.422% | 12.84% | 19.27% | 628.5634 | 694.728
Pr, ppm | 5.0923 | 0.60772 | 3.8769 | 6.3078 | 3.2692 | 6.9155 | 11.93% | 23.87% | 35.8% | 4.8377 | 5.3469
Rb, ppm | 65.2767 | 3.5841 | 58.1085 | 72.445 | 54.5244 | 76.0291 | 5.491% | 10.98% | 16.47% | 62.0129 | 68.5406
S, wt.% | 4.0693 | 0.14605 | 3.7772 | 4.3614 | 3.6312 | 4.5075 | 3.589% | 7.178% | 10.77% | 3.8659 | 4.2728
Sb, ppm | 298.8757 | 16.4481 | 265.9795 | 331.7719 | 249.5314 | 348.22 | 5.503% | 11.01% | 16.51% | 283.9319 | 313.8195
Sc, ppm | 3.0648 | 0.23996 | 2.5849 | 3.5447 | 2.3449 | 3.7847 | 7.83% | 15.66% | 23.49% | 2.9116 | 3.218
Se, ppm | 29.123 | 2.9742 | 23.1746 | 35.0713 | 20.2005 | 38.0455 | 10.21% | 20.42% | 30.64% | 27.6668 | 30.5791
Sm, ppm | 3.6759 | 0.28095 | 3.114 | 4.2378 | 2.833 | 4.5187 | 7.643% | 15.29% | 22.93% | 3.4921 | 3.8597
Sn, ppm | 27.1181 | 2.0177 | 23.0826 | 31.1536 | 21.0648 | 33.1713 | 7.441% | 14.88% | 22.32% | 25.7622 | 28.474
Sr, ppm | 305.75 | 27.7162 | 250.3176 | 361.1824 | 222.6014 | 388.8986 | 9.065% | 18.13% | 27.19% | 290.4625 | 321.0375
Ta, ppm | 0.67739 | 0.042214 | 0.59296 | 0.76182 | 0.55075 | 0.80403 | 6.232% | 12.46% | 18.7% | 0.64352 | 0.71126
Tb, ppm | 0.30958 | 0.061331 | 0.18692 | 0.43225 | 0.12559 | 0.49358 | 19.81% | 39.62% | 59.43% | 0.2941 | 0.32506
Te, ppm | 41.5506 | 2.2113 | 37.128 | 45.9732 | 34.9167 | 48.1845 | 5.322% | 10.64% | 15.97% | 39.4731 | 43.6281
Th, ppm | 8.9466 | 1.0658 | 6.815 | 11.0782 | 5.7492 | 12.1441 | 11.91% | 23.83% | 35.74% | 8.4993 | 9.394
Ti, wt.% | 0.16664 | 0.0056041 | 0.15543 | 0.17784 | 0.14982 | 0.18345 | 3.363% | 6.726% | 10.09% | 0.1583 | 0.17497
Tl, ppm | 1.8371 | 0.1124 | 1.6123 | 2.0619 | 1.4999 | 2.1743 | 6.118% | 12.24% | 18.36% | 1.7452 | 1.9289
U, ppm | 2.6326 | 0.11591 | 2.4008 | 2.8644 | 2.2849 | 2.9803 | 4.403% | 8.806% | 13.21% | 2.501 | 2.7642
V, ppm | 30.4801 | 1.3166 | 27.8469 | 33.1132 | 26.5303 | 34.4298 | 4.319% | 8.639% | 12.96% | 28.9561 | 32.0041
W, ppm | 7.5661 | 0.41993 | 6.7263 | 8.406 | 6.3063 | 8.8259 | 5.55% | 11.1% | 16.65% | 7.1878 | 7.9444
Y, ppm | 6.5911 | 0.36598 | 5.8591 | 7.3231 | 5.4932 | 7.6891 | 5.553% | 11.11% | 16.66% | 6.2616 | 6.9207
Yb, ppm | 0.51952 | 0.053051 | 0.41342 | 0.62563 | 0.36037 | 0.67868 | 10.21% | 20.42% | 30.63% | 0.49355 | 0.5455
Zn, ppm | 1753.6095 | 74.4032 | 1604.8031 | 1902.416 | 1530.3999 | 1976.8192 | 4.243% | 8.486% | 12.73% | 1665.929 | 1841.29
Zr, ppm | 60.6013 | 5.2036 | 50.1941 | 71.0084 | 44.9905 | 76.212 | 8.587% | 17.17% | 25.76% | 57.5712 | 63.6313
Aqua Regia Digestion
Ag, ppm | 48.4286 | 2.0212 | 44.3863 | 52.4709 | 42.3651 | 54.492 | 4.173% | 8.347% | 12.52% | 46.0071 | 50.85
Al, wt.% | 0.84724 | 0.057719 | 0.7318 | 0.96268 | 0.67408 | 1.0204 | 6.813% | 13.63% | 20.44% | 0.80488 | 0.8896
As, ppm | 2806.8049 | 151.3001 | 2504.2047 | 3109.4051 | 2352.9046 | 3260.7052 | 5.39% | 10.78% | 16.17% | 2666.4647 | 2947.1451
Be, ppm | 0.28972 | 0.023963 | 0.2418 | 0.33765 | 0.21783 | 0.36161 | 8.271% | 16.54% | 24.81% | 0.27524 | 0.30421
Bi, ppm | 219.8074 | 11.3388 | 197.1298 | 242.4851 | 185.7909 | 253.8239 | 5.159% | 10.32% | 15.48% | 208.817 | 230.7978
Ca, wt.% | 0.11989 | 0.0062078 | 0.10748 | 0.13231 | 0.10127 | 0.13852 | 5.178% | 10.36% | 15.53% | 0.1139 | 0.12589
Cd, ppm | 12.253 | 0.58744 | 11.0781 | 13.4279 | 10.4907 | 14.0153 | 4.794% | 9.589% | 14.38% | 11.6403 | 12.8656
Ce, ppm | 13.7282 | 1.0618 | 11.6045 | 15.8519 | 10.5427 | 16.9137 | 7.735% | 15.47% | 23.2% | 13.0418 | 14.4146
Co, ppm | 7.7258 | 0.37226 | 6.9813 | 8.4703 | 6.609 | 8.8426 | 4.818% | 9.637% | 14.46% | 7.3395 | 8.1121
Cr, ppm | 33.0658 | 2.5741 | 27.9175 | 38.214 | 25.3434 | 40.7881 | 7.785% | 15.57% | 23.35% | 31.4125 | 34.7191
Cs, ppm | 0.73729 | 0.032448 | 0.6724 | 0.80219 | 0.63995 | 0.83463 | 4.401% | 8.802% | 13.2% | 0.70043 | 0.77416
Cu, wt.% | 0.97152 | 0.026755 | 0.91801 | 1.025 | 0.89126 | 1.0518 | 2.754% | 5.508% | 8.262% | 0.92295 | 1.0201
Fe, wt.% | 2.2728 | 0.12237 | 2.028 | 2.5175 | 1.9057 | 2.6399 | 5.384% | 10.77% | 16.15% | 2.1591 | 2.3864
Ga, ppm | 6.3593 | 0.42438 | 5.5105 | 7.208 | 5.0861 | 7.6324 | 6.673% | 13.35% | 20.02% | 6.0413 | 6.6772
Hf, ppm | 0.38357 | 0.028605 | 0.32636 | 0.44078 | 0.29775 | 0.46938 | 7.458% | 14.92% | 22.37% | 0.36439 | 0.40275
Hg, ppm | 0.80046 | 0.059215 | 0.68202 | 0.91889 | 0.62281 | 0.9781 | 7.398% | 14.8% | 22.19% | 0.76043 | 0.84048
In, ppm | 3.7628 | 0.10665 | 3.5495 | 3.9761 | 3.4428 | 4.0828 | 2.834% | 5.669% | 8.503% | 3.5747 | 3.9509
K, wt.% | 0.21343 | 0.01843 | 0.17657 | 0.25029 | 0.15814 | 0.26872 | 8.635% | 17.27% | 25.9% | 0.20276 | 0.2241
La, ppm | 6.6766 | 0.48237 | 5.7119 | 7.6414 | 5.2295 | 8.1238 | 7.225% | 14.45% | 21.67% | 6.3428 | 7.0105
Li, ppm | 8.4614 | 0.92896 | 6.6035 | 10.3193 | 5.6745 | 11.2482 | 10.98% | 21.96% | 32.94% | 8.0383 | 8.8844
Mg, ppm | 1058.8625 | 72.8538 | 913.1548 | 1204.5702 | 840.301 | 1277.424 | 6.88% | 13.76% | 20.64% | 1005.9194 | 1111.8056
Mn, ppm | 65.5196 | 3.0789 | 59.3618 | 71.6774 | 56.2828 | 74.7563 | 4.699% | 9.398% | 14.1% | 62.2436 | 68.7956
Mo, ppm | 4.4737 | 0.30302 | 3.8677 | 5.0798 | 3.5647 | 5.3828 | 6.773% | 13.55% | 20.32% | 4.25 | 4.6974
Na, wt.% | 0.048667 | 0.0098371 | 0.028993 | 0.068341 | 0.019156 | 0.078178 | 20.21% | 40.43% | 60.64% | 0.046234 | 0.0511
Nb, ppm | 0.16208 | 0.027286 | 0.10751 | 0.21665 | 0.080227 | 0.24394 | 16.83% | 33.67% | 50.5% | 0.15398 | 0.17019
Ni, ppm | 24.2843 | 1.4335 | 21.4174 | 27.1513 | 19.9839 | 28.5848 | 5.903% | 11.81% | 17.71% | 23.0701 | 25.4985
P, ppm | 249.4704 | 19.2273 | 211.0159 | 287.9249 | 191.7886 | 307.1522 | 7.707% | 15.41% | 23.12% | 236.9969 | 261.9439
Pb, ppm | 511.6994 | 20.5269 | 470.6457 | 552.7531 | 450.1188 | 573.28 | 4.012% | 8.023% | 12.03% | 486.1144 | 537.2844
Rb, ppm | 7.6278 | 0.60925 | 6.4093 | 8.8463 | 5.8 | 9.4555 | 7.987% | 15.97% | 23.96% | 7.2464 | 8.0092
S, wt.% | 2.6502 | 0.10873 | 2.4328 | 2.8677 | 2.3241 | 2.9764 | 4.103% | 8.205% | 12.31% | 2.5177 | 2.7827
Sb, ppm | 264.7681 | 14.0059 | 236.7562 | 292.78 | 222.7502 | 306.7859 | 5.29% | 10.58% | 15.87% | 251.5297 | 278.0065
Sc, ppm | 0.84412 | 0.09555 | 0.65301 | 1.0352 | 0.55746 | 1.1308 | 11.32% | 22.64% | 33.96% | 0.80191 | 0.88632
Se, ppm | 27.7179 | 3.57 | 20.5778 | 34.8579 | 17.0078 | 38.4279 | 12.88% | 25.76% | 38.64% | 26.332 | 29.1038
Sn, ppm | 24.7736 | 1.2198 | 22.3339 | 27.2132 | 21.1141 | 28.4331 | 4.924% | 9.848% | 14.77% | 23.5349 | 26.0123
Sr, ppm | 38.6366 | 6.2529 | 26.1308 | 51.1425 | 19.8779 | 57.3954 | 16.18% | 32.37% | 48.55% | 36.7048 | 40.5685
Te, ppm | 41.7054 | 2.3463 | 37.0128 | 46.3979 | 34.6665 | 48.7442 | 5.626% | 11.25% | 16.88% | 39.6201 | 43.7906
Th, ppm | 3.0845 | 0.22815 | 2.6283 | 3.5408 | 2.4001 | 3.769 | 7.396% | 14.79% | 22.19% | 2.9303 | 3.2388
Tl, ppm | 1.4893 | 0.052485 | 1.3843 | 1.5943 | 1.3318 | 1.6467 | 3.524% | 7.048% | 10.57% | 1.4148 | 1.5637
U, ppm | 1.1203 | 0.078311 | 0.96364 | 1.2769 | 0.88533 | 1.3552 | 6.99% | 13.98% | 20.97% | 1.0642 | 1.1763
V, ppm | 11.5915 | 0.84161 | 9.9082 | 13.2747 | 9.0666 | 14.1163 | 7.261% | 14.52% | 21.78% | 11.0119 | 12.171
W, ppm | 3.5799 | 0.4236 | 2.7327 | 4.4271 | 2.3091 | 4.8507 | 11.83% | 23.67% | 35.5% | 3.4009 | 3.7589
Y, ppm | 3.0932 | 0.16923 | 2.7547 | 3.4316 | 2.5855 | 3.6009 | 5.471% | 10.94% | 16.41% | 2.9385 | 3.2478
Zn, ppm | 1763.5635 | 61.9004 | 1639.7626 | 1887.3644 | 1577.8622 | 1949.2649 | 3.51% | 7.02% | 10.53% | 1675.3854 | 1851.7417
Zr, ppm | 11.0958 | 0.85204 | 9.3917 | 12.7999 | 8.5397 | 13.6519 | 7.679% | 15.36% | 23.04% | 10.541 | 11.6506

## Confidence & Tolerance Limits
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 610
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 9.8341 | 0.25443 | 9.7338 | 9.9343 | 9.7945 | 9.8737
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 9.8119 | 0.34094 | 9.6146 | 10.0092 | 9.7686 | 9.8551
Infrared Combustion
S, Sulphur (wt.%) | 4.1172 | 0.15678 | 4.0473 | 4.1871 | 4.068 | 4.1664
4-Acid Digestion
Ag, Silver (ppm) | 49.3686 | 1.7899 | 48.655 | 50.0822 | 48.0891 | 50.6481
Al, Aluminium (wt.%) | 5.9878 | 0.26049 | 5.874 | 6.1016 | 5.8422 | 6.1335
As, Arsenic (ppm) | 2834.5718 | 185.1685 | 2750.786 | 2918.3575 | 2781.1935 | 2887.9501
Be, Beryllium (ppm) | 1.4751 | 0.10935 | 1.4259 | 1.5242 | 1.3997 | 1.5504
Bi, Bismuth (ppm) | 223.5764 | 9.6809 | 219.2321 | 227.9206 | 218.8826 | 228.2702
Ca, Calcium (wt.%) | 0.24074 | 0.014085 | 0.23456 | 0.24693 | 0.23179 | 0.2497
Cd, Cadmium (ppm) | 12.0867 | 0.70848 | 11.7647 | 12.4087 | 11.7455 | 12.428
Ce, Cerium (ppm) | 46.6884 | 4.4632 | 44.1336 | 49.2433 | 44.877 | 48.4999
Co, Cobalt (ppm) | 7.7238 | 0.30497 | 7.6049 | 7.8426 | 7.4283 | 8.0192
Cr, Chromium (ppm) | 39.1331 | 5.0643 | 37.1609 | 41.1053 | 37.1407 | 41.1254
Cs, Caesium (ppm) | 2.1551 | 0.14615 | 2.0851 | 2.2251 | 2.0612 | 2.249
Cu, Copper (wt.%) | 0.97124 | 0.022851 | 0.96185 | 0.98063 | 0.95759 | 0.9849
Dy, Dysprosium (ppm) | 1.525 | 0.08893 | 1.4136 | 1.6364 | IND | IND
Er, Erbium (ppm) | 0.59367 | 0.040994 | 0.56146 | 0.62587 | IND | IND
Eu, Europium (ppm) | 0.77133 | 0.050769 | 0.72603 | 0.81663 | IND | IND
Fe, Iron (wt.%) | 2.3658 | 0.076692 | 2.3341 | 2.3976 | 2.3195 | 2.4122
Ga, Gallium (ppm) | 24.419 | 1.0235 | 23.9371 | 24.901 | 23.4105 | 25.4275
Gd, Gadolinium (ppm) | 2.856 | 0.25636 | 2.5304 | 3.1816 | 2.6567 | 3.0553
Hf, Hafnium (ppm) | 2.0068 | 0.12755 | 1.945 | 2.0687 | 1.9338 | 2.0798
Ho, Holmium (ppm) | 0.217 | 0.021359 | 0.18801 | 0.24599 | IND | IND
In, Indium (ppm) | 3.8952 | 0.24442 | 3.7733 | 4.0171 | 3.7845 | 4.0059
K, Potassium (wt.%) | 1.9786 | 0.067443 | 1.9509 | 2.0064 | 1.9447 | 2.0125
La, Lanthanum (ppm) | 20.3844 | 4.4167 | 18.3041 | 22.4647 | 19.4461 | 21.3227
Li, Lithium (ppm) | 28.4423 | 1.3454 | 27.9118 | 28.9728 | 27.2597 | 29.6249
Mg, Magnesium (ppm) | 1557.4306 | 90.5031 | 1520.1821 | 1594.6791 | 1514.1914 | 1600.6698
Mn, Manganese (ppm) | 77.7201 | 3.5584 | 76.3717 | 79.0685 | 75.4456 | 79.9946
Mo, Molybdenum (ppm) | 4.8165 | 0.33191 | 4.6774 | 4.9556 | 4.5833 | 5.0497
Na, Sodium (wt.%) | 0.82035 | 0.026705 | 0.80954 | 0.83116 | 0.80303 | 0.83767
Nb, Niobium (ppm) | 8.2736 | 0.42983 | 8.0491 | 8.498 | 8.0002 | 8.5469
Nd, Neodymium (ppm) | 18.881 | 1.5596 | 17.0068 | 20.7552 | 17.549 | 20.213
Ni, Nickel (ppm) | 24.6174 | 1.1675 | 24.1745 | 25.0602 | 23.7812 | 25.4535
P, Phosphorus (ppm) | 554.2234 | 36.413 | 537.569 | 570.8778 | 538.5797 | 569.8672
Pb, Lead (ppm) | 661.6457 | 42.4928 | 643.5506 | 679.7408 | 650.1256 | 673.1658
Pr, Praseodymium (ppm) | 5.0923 | 0.60772 | 4.296 | 5.8887 | 4.8058 | 5.3789
Rb, Rubidium (ppm) | 65.2767 | 3.5841 | 63.5465 | 67.007 | 63.3571 | 67.1964
S, Sulphur (wt.%) | 4.0693 | 0.14605 | 4.0095 | 4.1291 | 4.005 | 4.1337
Sb, Antimony (ppm) | 298.8757 | 16.4481 | 291.234 | 306.5174 | 291.0432 | 306.7082
Sc, Scandium (ppm) | 3.0648 | 0.23996 | 2.9456 | 3.184 | 2.9159 | 3.2137
Se, Selenium (ppm) | 29.123 | 2.9742 | 27.6338 | 30.6121 | 27.6637 | 30.5822
Sm, Samarium (ppm) | 3.6759 | 0.28095 | 3.352 | 3.9998 | 3.509 | 3.8428
Sn, Tin (ppm) | 27.1181 | 2.0177 | 26.272 | 27.9641 | 26.1449 | 28.0913
Sr, Strontium (ppm) | 305.75 | 27.7162 | 292.7515 | 318.7485 | 297.8261 | 313.6739
Ta, Tantalum (ppm) | 0.67739 | 0.042214 | 0.65219 | 0.70259 | 0.63953 | 0.71525
Tb, Terbium (ppm) | 0.30958 | 0.061331 | 0.25684 | 0.36233 | IND | IND
Te, Tellurium (ppm) | 41.5506 | 2.2113 | 40.4424 | 42.6588 | 39.7383 | 43.3629
Th, Thorium (ppm) | 8.9466 | 1.0658 | 8.3625 | 9.5308 | 8.5723 | 9.3209
Ti, Titanium (wt.%) | 0.16664 | 0.0056041 | 0.16437 | 0.1689 | 0.16267 | 0.1706
Tl, Thallium (ppm) | 1.8371 | 0.1124 | 1.7797 | 1.8944 | 1.778 | 1.8961
U, Uranium (ppm) | 2.6326 | 0.11591 | 2.5776 | 2.6876 | 2.5318 | 2.7334
V, Vanadium (ppm) | 30.4801 | 1.3166 | 29.9713 | 30.9889 | 29.5829 | 31.3773
W, Tungsten (ppm) | 7.5661 | 0.41993 | 7.3543 | 7.7779 | 7.2815 | 7.8507
Y, Yttrium (ppm) | 6.5911 | 0.36598 | 6.4136 | 6.7686 | 6.4049 | 6.7773
Yb, Ytterbium (ppm) | 0.51952 | 0.053051 | 0.48072 | 0.55833 | IND | IND
Zn, Zinc (ppm) | 1753.6095 | 74.4032 | 1721.8071 | 1785.4119 | 1717.1912 | 1790.0279
Zr, Zirconium (ppm) | 60.6013 | 5.2036 | 58.4135 | 62.789 | 58.6925 | 62.51
Aqua Regia Digestion
Ag, Silver (ppm) | 48.4286 | 2.0212 | 47.5528 | 49.3044 | 47.5039 | 49.3533
Al, Aluminium (wt.%) | 0.84724 | 0.057719 | 0.82004 | 0.87444 | 0.82619 | 0.86829
As, Arsenic (ppm) | 2806.8049 | 151.3001 | 2739.8602 | 2873.7496 | 2756.9868 | 2856.623
Be, Beryllium (ppm) | 0.28972 | 0.023963 | 0.2846 | 0.29484 | 0.25275 | 0.32669
Bi, Bismuth (ppm) | 219.8074 | 11.3388 | 214.6025 | 225.0123 | 215.665 | 223.9499
Ca, Calcium (wt.%) | 0.11989 | 0.0062078 | 0.11731 | 0.12247 | 0.11604 | 0.12374
Cd, Cadmium (ppm) | 12.253 | 0.58744 | 11.9706 | 12.5354 | 12.0409 | 12.4651
Ce, Cerium (ppm) | 13.7282 | 1.0618 | 13.1005 | 14.3559 | 13.2075 | 14.2489
Co, Cobalt (ppm) | 7.7258 | 0.37226 | 7.5532 | 7.8984 | 7.4614 | 7.9903
Cr, Chromium (ppm) | 33.0658 | 2.5741 | 31.94 | 34.1915 | 31.9693 | 34.1622
Cs, Caesium (ppm) | 0.73729 | 0.032448 | 0.71766 | 0.75693 | 0.7115 | 0.76308
Cu, Copper (wt.%) | 0.97152 | 0.026755 | 0.96046 | 0.98259 | 0.95698 | 0.98607
Fe, Iron (wt.%) | 2.2728 | 0.12237 | 2.2214 | 2.3242 | 2.2318 | 2.3138
Ga, Gallium (ppm) | 6.3593 | 0.42438 | 6.1703 | 6.5482 | 6.105 | 6.6135
Hf, Hafnium (ppm) | 0.38357 | 0.028605 | 0.36808 | 0.39906 | 0.36492 | 0.40221
Hg, Mercury (ppm) | 0.80046 | 0.059215 | 0.76378 | 0.83713 | 0.77054 | 0.83037
In, Indium (ppm) | 3.7628 | 0.10665 | 3.711 | 3.8146 | 3.6421 | 3.8835
K, Potassium (wt.%) | 0.21343 | 0.01843 | 0.20508 | 0.22179 | 0.20411 | 0.22275
La, Lanthanum (ppm) | 6.6766 | 0.48237 | 6.436 | 6.9173 | 6.473 | 6.8803
Li, Lithium (ppm) | 8.4614 | 0.92896 | 7.9692 | 8.9535 | 8.0748 | 8.848
Mg, Magnesium (ppm) | 1058.8625 | 72.8538 | 1029.3338 | 1088.3912 | 1023.5127 | 1094.2123
Mn, Manganese (ppm) | 65.5196 | 3.0789 | 64.2548 | 66.7844 | 63.9673 | 67.0718
Mo, Molybdenum (ppm) | 4.4737 | 0.30302 | 4.3421 | 4.6053 | 4.2157 | 4.7317
Na, Sodium (wt.%) | 0.048667 | 0.0098371 | 0.04425 | 0.053084 | 0.046423 | 0.050911
Nb, Niobium (ppm) | 0.16208 | 0.027286 | 0.13902 | 0.18514 | IND | IND
Ni, Nickel (ppm) | 24.2843 | 1.4335 | 23.6897 | 24.879 | 23.3631 | 25.2056
P, Phosphorus (ppm) | 249.4704 | 19.2273 | 239.5007 | 259.4402 | 240.5753 | 258.3655
Pb, Lead (ppm) | 511.6994 | 20.5269 | 502.9622 | 520.4366 | 502.4275 | 520.9713
Rb, Rubidium (ppm) | 7.6278 | 0.60925 | 7.2485 | 8.0071 | 7.3965 | 7.8591
S, Sulphur (wt.%) | 2.6502 | 0.10873 | 2.6012 | 2.6993 | 2.6014 | 2.6991
Sb, Antimony (ppm) | 264.7681 | 14.0059 | 257.966 | 271.5702 | 259.6098 | 269.9264
Sc, Scandium (ppm) | 0.84412 | 0.09555 | 0.79492 | 0.89331 | IND | IND
Se, Selenium (ppm) | 27.7179 | 3.57 | 25.8685 | 29.5672 | 26.8865 | 28.5493
Sn, Tin (ppm) | 24.7736 | 1.2198 | 24.2177 | 25.3295 | 24.0718 | 25.4753
Sr, Strontium (ppm) | 38.6366 | 6.2529 | 35.7599 | 41.5133 | 37.2653 | 40.008
Te, Tellurium (ppm) | 41.7054 | 2.3463 | 40.5842 | 42.8265 | 40.5875 | 42.8233
Th, Thorium (ppm) | 3.0845 | 0.22815 | 2.9612 | 3.2079 | 2.965 | 3.2041
Tl, Thallium (ppm) | 1.4893 | 0.052485 | 1.4636 | 1.5149 | 1.4428 | 1.5357
U, Uranium (ppm) | 1.1203 | 0.078311 | 1.0803 | 1.1602 | 1.077 | 1.1635
V, Vanadium (ppm) | 11.5915 | 0.84161 | 11.2042 | 11.9788 | 11.1589 | 12.024
W, Tungsten (ppm) | 3.5799 | 0.4236 | 3.3515 | 3.8083 | 3.4423 | 3.7175
Y, Yttrium (ppm) | 3.0932 | 0.16923 | 3.0097 | 3.1766 | 3.0004 | 3.186
Zn, Zinc (ppm) | 1763.5635 | 61.9004 | 1737.869 | 1789.2581 | 1733.344 | 1793.783
Zr, Zirconium (ppm) | 11.0958 | 0.85204 | 10.692 | 11.4995 | 10.6655 | 11.526

## Indicative Values
Table 3. Indicative Values for OREAS 610
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 5 | Pt | ppb | < 5
Infrared Combustion
C | wt.% | 0.07
4-Acid Digestion
Ba | ppm | 242.4673 | Hg | ppm | 0.42833 | Re | ppb | < 2
Ge | ppm | 2.1336 | Lu | ppb | 80 | Tm | ppb | 91.6667
Aqua Regia Digestion
B | ppm | < 10 | Ho | ppm | 0.091667 | Sm | ppm | 1.55
Ba | ppm | 83.9675 | Lu | ppb | 31.6667 | Ta | ppm | < 0.01
Dy | ppm | 0.81667 | Nd | ppm | 7.885 | Tb | ppm | 0.165
Er | ppm | 0.26667 | Pd | ppb | 173.6667 | Ti | wt.% | 0.0081878
Eu | ppm | 0.23333 | Pr | ppm | 1.9333 | Tm | ppb | < 100
Gd | ppm | 1.4333 | Pt | ppb | < 2 | Yb | ppm | 0.2
Ge | ppm | 0.28419 | Re | ppb | 1.1042
Borate Fusion XRF
Al2O3 | wt.% | 11.72 | Fe2O3 | wt.% | 3.385 | S | wt.% | 4.2047
As | ppm | 3035 | K2O | wt.% | 2.405 | SiO2 | wt.% | 69.21
BaO | ppm | 4615 | MgO | wt.% | 0.29 | Sn | ppm | 10
CaO | wt.% | 0.33 | MnO | wt.% | 0.0105 | Sr | ppm | 338.2378
Cl | ppm | 35 | Na2O | wt.% | 1.15 | TiO2 | wt.% | 0.278
Co | ppm | 15 | Ni | ppm | 35 | V2O5 | ppm | 65
Cr2O3 | ppm | 65 | P2O5 | wt.% | 0.125 | Zn | ppm | 1795
Cu | wt.% | 0.9685 | Pb | ppm | 710 | Zr | ppm | 144.3628
Thermogravimetry
LOI1000 | wt.% | 8.74
Laser Ablation ICP-MS
Ag | ppm | 57.2 | Hf | ppm | 3.82 | Sm | ppm | 3.94
As | ppm | 2980 | Ho | ppm | 0.25 | Sn | ppm | 26.4
Ba | ppm | 3865 | In | ppm | 3.525 | Sr | ppm | 333
Be | ppm | 1.8 | La | ppm | 26.15 | Ta | ppm | 0.73
Bi | ppm | 220 | Lu | ppb | 85 | Tb | ppm | 0.36
Cd | ppm | 14 | Mn | ppm | 72.5 | Te | ppm | 45.3
Ce | ppm | 49.45 | Mo | ppm | 4.6 | Th | ppm | 10.05
Co | ppm | 7.45 | Nb | ppm | 8.365 | Ti | wt.% | 0.164
Cr | ppm | 46.5 | Nd | ppm | 20.05 | Tl | ppm | 2.7
Cs | ppm | 2.115 | Ni | ppm | 29 | Tm | ppb | 105
Cu | wt.% | 0.973 | Pb | ppm | 731.5 | U | ppm | 2.63
Dy | ppm | 1.54 | Pr | ppm | 5.445 | V | ppm | 30.95
Er | ppm | 0.6 | Rb | ppm | 62 | W | ppm | 7.25
Eu | ppm | 0.78 | Re | ppb | < 10 | Y | ppm | 7.11
Ga | ppm | 23.1 | Sb | ppm | 312 | Yb | ppm | 0.65
Gd | ppm | 2.785 | Sc | ppm | 3.55 | Zn | ppm | 1765
Ge | ppm | 4.55 | Se | ppm | < 5 | Zr | ppm | 125.5
X-ray Photon Assay
Au | ppm | 9.8665

## Fire Assay
| Analytical results for Au in OREAS 610 (Certified Value 9.83 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*GRAV | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*GRAV | 2
|  |  | 0.085g | 25g | 25g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 40g | 30g | 3
| 1 | 1 | 10.1424 | 9.42 | 10 | 9.813 | 9.42 | 9.5 | 10.6 | 10.2 | 10 | 9.85 | 10.1 | 9.351 | 9.9258 | 10.1 | 10.457 | 9.26 | 9.9 | 9.8327 | 9.983 | 9.82 | 9.423 | 9.96 | 9.69 | 9.22 | 9.89 | 10.68 | 1
| 1 | 2 | 9.7649 | 9.41 | 10.09 | 9.613 | 9.55 | 9.6 | 10.2 | 10 | 10.05 | 9.86 | 10 | 9.752 | 9.98 | 10 | 10.226 | 9.51 | 9.87 | 9.656 | 9.914 | 9.78 | 9.525 | 10.27 | 9.49 | 9.12 | 9.86 | 10 | 45
| 1 | 3 | 10.2033 | 9.53 | 10.012 | 9.643 | 9.79 | 9.56 | 10.55 | 10.15 | 9.61 | 9.7 | 9.71 | 9.449 | 9.915 | 10.1 | 10.233 | 9.62 | 9.71 | 9.6933 | 9.875 | 9.67 | 9.464 | 10.09 | 9.64 | 8.98 | 9.93 | 10.43 | 16
| 1 | 4 | 10.522 | 9.33 | 10.06 | 9.814 | 9.83 | 9.56 | 10.7 | 9.91 | 9.94 | 9.93 | 10.1 | 9.542 | 9.93 | 10 | 10.124 | 9.96 | 9.68 | 9.6787 | 9.919 | 9.65 | 9.662 | 10.24 | 9.76 | 8.81 | 9.88 | 10.02 | 9.8341
| 1 | 5 | 9.9813 | 9.41 | 9.943 | 9.201 | 9.75 | 9.68 | 10.65 | 10 | 10.2 | 9.88 | 10.2 | 9.629 | 9.965 | 9.8 | 10.397 | 9.88 | 9.86 | 9.6887 | 9.916 | 9.95 | 9.57 | 10.03 | 9.5 | 8.98 | 9.91 | 10.0895 | 7
| 1 | 6 | 10.1929 | 9.3 | 10.068 | 9.484 | 9.96 | 9.6 | 10.55 | 9.95 | 10.1 | 9.93 | 9.92 | 9.487 | 9.945 | 10 | 10.168 | 9.59 | 9.86 | 9.6973 | 9.949 | 9.88 | 9.514 | 10.14 | 9.6 | 8.88 | 9.85 | 10.74
|  | 7 | 00: 10.172
|  | 8 | 00: 10.5475
|  | 9 | 00: 9.9789
|  | 10 | 00: 10.2849
|  | 11 | 00: 9.8752
|  | 12 | 00: 9.8439
|  | 13 | 00: 9.8782
|  | 14 | 00: 10.249
|  | 15 | 00: 10.1568
|  | 16 | 00: 9.8613
|  | 17 | 00: 10.294
|  | 18 | 00: 9.8498
|  | 19 | 00: 9.9791
|  | 20 | 00: 9.6866
| Mean |  | 10.0732 | 9.4 | 10.0288 | 9.5947 | 9.7167 | 9.5833 | 10.5417 | 10.035 | 9.9833 | 9.8583 | 10.005 | 9.535 | 9.9435 | 10 | 10.2675 | 9.6367 | 9.8133 | 9.7078 | 9.926 | 9.7917 | 9.5263 | 10.1217 | 9.6133 | 8.9983 | 9.8867 | 10.3266
| Median |  | 10.0618 | 9.41 | 10.036 | 9.628 | 9.77 | 9.58 | 10.575 | 10 | 10.025 | 9.87 | 10.05 | 9.5145 | 9.9375 | 10 | 10.2295 | 9.605 | 9.86 | 9.691 | 9.9175 | 9.8 | 9.5195 | 10.115 | 9.62 | 8.98 | 9.885 | 10.2598
| Std Dev. |  | 0.24049 | 0.080498 | 0.054275 | 0.23051 | 0.19715 | 0.059889 | 0.17725 | 0.11467 | 0.20324 | 0.084715 | 0.17341 | 0.14109 | 0.024913 | 0.10954 | 0.13121 | 0.25461 | 0.09331 | 0.062942 | 0.036529 | 0.1172 | 0.083701 | 0.1199 | 0.10614 | 0.15105 | 0.030111 | 0.33567
| Rel.Std.Dev. |  | 2.387% | 0.8564% | 0.5412% | 2.402% | 2.029% | 0.6249% | 1.681% | 1.143% | 2.036% | 0.8593% | 1.733% | 1.48% | 0.2505% | 1.095% | 1.278% | 2.642% | 0.9508% | 0.6484% | 0.368% | 1.197% | 0.8786% | 1.185% | 1.104% | 1.679% | 0.3046% | 3.251%
| PDM3 |  | 2.431% | -4.414% | 1.98% | -2.435% | -1.194% | -2.55% | 7.195% | 2.043% | 1.518% | 0.2465% | 1.738% | -3.041% | 1.112% | 1.687% | 4.407% | -2.008% | -0.2111% | -1.284% | 0.9346% | -0.4314% | -3.13% | 2.924% | -2.245% | -8.499% | 0.5346% | 5.008%
| Z-Score (Absolute) |  | NA | 1.75 | 0.65 | 1.01 | 0.54 | 1.05 | 2.61 | 0.67 | 0.48 | 0 | 0.56 | 1.23 | 0.32 | 0.54 | 1.56 | 0.85 | 0.17 | 0.57 | 0.26 | 0.25 | 1.27 | 1.01 | 0.94 | 3.28 | 0.11 | 1.79
| Analytical results for Pd in OREAS 610 (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 00: Lab | 1
| No | No | 00: 2.16 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 40g | 2
| 1 | 1 | 00: < 5 | 1
| 1 | 2 | 00: < 5 | 1
| 1 | 3 | 00: < 5 | 16
| 1 | 4 | 00: < 5 | < 5
| 1 | 5 | 00: < 5 | 7
| 1 | 6 | 00: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00
| Analytical results for Pt in OREAS 610 (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 00: Lab | 1
| No | No | 00: 2.16 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 40g | 2
| 1 | 1 | 00: < 5 | 1
| 1 | 2 | 00: < 5 | 1
| 1 | 3 | 00: < 5 | 16
| 1 | 4 | 00: < 5 | < 5
| 1 | 5 | 00: < 5 | 7
| 1 | 6 | 00: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 610 (Certified Value 9.81 ppm) | [16 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.16 | 2.24 | 2.25 | ppm
|  |  | INAA | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | 2
|  |  | 0.085g | 15g | 30g | 25g | 25g | 25g | 25g | 25g | 30g | 25g | 25g | 15g | 40g | 30g | 25g | 3
| 1 | 1 | 10.1424 | 10.7854 | 9.71 | 10.15 | 10.4 | 10.15 | 9.71 | 9.35 | 9.59 | 9.34 | 9.8531 | > 10 | 9.83 | 9.51 | 9.86 | 1
| 1 | 2 | 9.7649 | 10.5184 | 9.95 | 9.78 | 10.3 | 10.05 | 9.36 | 9.35 | 9.6 | 9.5 | 9.8046 | > 10 | 9.83 | 9.49 | 9.83 | 45
| 1 | 3 | 10.2033 | 10.4833 | 9.74 | 10.05 | 10 | 10.3 | 9.51 | 9.49 | 9.6 | 9.53 | 9.6671 | > 10 | 9.47 | 9.48 | 9.8 | 16
| 1 | 4 | 10.522 | 10.4498 | 9.91 | 10.15 | 9.79 | 10.15 | 9.78 | 9.39 | 9.69 | 9.49 | 9.7521 | > 10 | 9.8 | 9.47 | 9.877 | 9.8119
| 1 | 5 | 9.9813 | 10.4995 | 9.74 | 9.96 | 10.05 | 10 | 9.81 | 9.15 | 9.69 | 9.59 | 9.6832 | > 10 | 9.38 | 9.42 | 9.84 | 9
| 1 | 6 | 10.1929 | 10.3425 | 9.96 | 9.98 | 10.15 | 10.2 | 9.91 | 9.13 | 9.8 | 9.23 | 9.9196 | > 10 | 9.83 | 9.46 | 9.81
|  | 7 | 00: 10.172
|  | 8 | 00: 10.5475
|  | 9 | 00: 9.9789
|  | 10 | 00: 10.2849
|  | 11 | 00: 9.8752
|  | 12 | 00: 9.8439
|  | 13 | 00: 9.8782
|  | 14 | 00: 10.249
|  | 15 | 00: 10.1568
|  | 16 | 00: 9.8613
|  | 17 | 00: 10.294
|  | 18 | 00: 9.8498
|  | 19 | 00: 9.9791
|  | 20 | 00: 9.6866
| Mean |  | 10.0732 | 10.5131 | 9.835 | 10.0117 | 10.115 | 10.1417 | 9.68 | 9.31 | 9.6617 | 9.4467 | 9.7799 |  | 9.69 | 9.4717 | 9.8362
| Median |  | 10.0618 | 10.4914 | 9.825 | 10.015 | 10.1 | 10.15 | 9.745 | 9.35 | 9.645 | 9.495 | 9.7783 |  | 9.815 | 9.475 | 9.835
| Std Dev. |  | 0.24049 | 0.14721 | 0.11675 | 0.13934 | 0.21897 | 0.10685 | 0.20591 | 0.14142 | 0.081833 | 0.13456 | 0.098345 |  | 0.20756 | 0.030605 | 0.02926
| Rel.Std.Dev. |  | 2.387% | 1.4% | 1.187% | 1.392% | 2.165% | 1.054% | 2.127% | 1.519% | 0.847% | 1.424% | 1.006% |  | 2.142% | 0.3231% | 0.2975%
| PDM3 |  | 2.663% | 7.147% | 0.2355% | 2.036% | 3.089% | 3.361% | -1.344% | -5.115% | -1.531% | -3.722% | -0.3256% |  | -1.242% | -3.467% | 0.2474%
| Z-Score (Absolute) |  | NA | 2.13 | 0.16 | 0.67 | 0.98 | 1.05 | 0.29 | 1.37 | 0.34 | 0.97 | 0 | NA | 0.26 | 0.9 | 0.16
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 610 (Indicative Value 0.07 wt.%) | [22 empty] | Indicative
C | Round | Replicate | BV Geo | [20 empty] | 1
| No | No | Perth | [20 empty] | wt.%
|  |  | IRC | [20 empty] | 3
3
| 1 | 1 | 0.07 | [20 empty] | 1
| 1 | 2 | 0.07 | [20 empty] | 3
| Mean |  | 0.07 | [20 empty] | 16
| Median |  | 0.07 | [20 empty] | 0.07
| Std Dev. |  | 0 | [20 empty] | 9
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 610 (Certified Value 4.12 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 3.99 | 3.82 | 4.17 | 4.3 | 4.25 | 4.14 | 4.12 | 4.18 | 4.05 | 4.19 | 4.3159 | 4.15 | 4.27 | 3.9 | 4.28 | 3.84 | 4.17 | 4.19 | 4.02 | 4.27 | 3.9305 | 1
| 1 | 2 | 4.07 | 3.96 | 4.03 | 4.33 | 4.28 | 4.08 | 4.11 | 4.28 | 4.04 | 4.27 | 4.3154 | 4.15 | 4.3 | 3.95 | 4.3 | 3.95 | 4.14 | 4.16 | 3.99 | 4.24 | 3.8971 | 16
| 1 | 3 | 4.06 | 3.9 | 4.15 | 4.29 | 4.32 | 4.08 | 4.18 | 4.2 | 4.03 | 4.23 | 4.3061 | 4.18 | 4.26 | 3.9 | 4.34 | 3.66 | 4.17 | 4.1 | 4.06 | 4.26 | 3.8625 | 16
| 1 | 4 | 4.14 | 3.79 | 4 | 4.34 | 4.2 | 4.04 | 4.1 | 4.14 | 4.03 | 4.15 | 4.3148 | 4.15 | 4.27 | 3.98 | 4.32 | 3.84 | 4.15 | 4.23 | 3.92 | 4.26 | 3.8535 | 4.1172
| 1 | 5 | 4.03 | 3.73 | 4 | 4.34 | 4.27 | 4.05 | 4.11 | 4.11 | 4.05 | 4.19 | 4.3236 | 4.11 | 4.32 | 3.94 | 4.34 | 3.78 | 4.11 | 4.21 | 3.96 | 4.21 | 3.8945 | 11
| 1 | 6 | 4 | 3.91 | 4.06 | 4.29 | 4.21 | 4.12 | 4.18 | 4.28 | 4.03 | 4.21 | 4.311 | 4.13 | 4.24 | 3.89 | 4.32 | 3.9 | 4.11 | 4.21 | 3.95 | 4.25 | 3.94
| Mean |  | 4.0483 | 3.8517 | 4.0683 | 4.315 | 4.255 | 4.085 | 4.1333 | 4.1983 | 4.0383 | 4.2067 | 4.3145 | 4.145 | 4.2767 | 3.9267 | 4.3167 | 3.8283 | 4.1417 | 4.1833 | 3.9833 | 4.2483 | 3.8964
| Median |  | 4.045 | 3.86 | 4.045 | 4.315 | 4.26 | 4.08 | 4.115 | 4.19 | 4.035 | 4.2 | 4.3151 | 4.15 | 4.27 | 3.92 | 4.32 | 3.84 | 4.145 | 4.2 | 3.975 | 4.255 | 3.8958
| Std Dev. |  | 0.054924 | 0.08612 | 0.074677 | 0.02429 | 0.045056 | 0.038859 | 0.036697 | 0.070545 | 0.0098319 | 0.040825 | 0.0058242 | 0.023452 | 0.028752 | 0.03559 | 0.023381 | 0.10088 | 0.027142 | 0.047188 | 0.050859 | 0.02137 | 0.034819
| Rel.Std.Dev. |  | 1.357% | 2.236% | 1.836% | 0.5629% | 1.059% | 0.9513% | 0.8878% | 1.68% | 0.2435% | 0.9705% | 0.135% | 0.5658% | 0.6723% | 0.9064% | 0.5416% | 2.635% | 0.6553% | 1.128% | 1.277% | 0.503% | 0.8936%
| PDM3 |  | -1.672% | -6.449% | -1.186% | 4.805% | 3.347% | -0.7817% | 0.3923% | 1.971% | -1.915% | 2.173% | 4.792% | 0.6756% | 3.874% | -4.627% | 4.845% | -7.016% | 0.5947% | 1.607% | -3.251% | 3.185% | -5.363%
| Z-Score (Absolute) |  | 0.59 | 1.83 | 0.46 | 1.1 | 0.72 | 0.36 | 0.05 | 0.36 | 0.65 | 0.41 | 1.09 | 0.02 | 0.85 | 1.36 | 1.11 | 1.98 | 0 | 0.26 | 1 | 0.67 | 1.55
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 610 (Certified Value 49.4 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*AAS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 49.3 | 49.296 | 47.6 | 47.4 | 48 | 47 | 48.7 | 50.2 | 50.5 | 51.24 | 50.829 | 48.83 | 46.2 | 49 | 49.9902 | 49.5 | 51 | > 10 | 49.77 | 47 | 46.6 | 51 | 48.44 | 1
| 1 | 2 | 48.5 | 48.999 | 47.2 | 48.4 | 48.4 | 49.5 | 48.4 | 51.5 | 50.1 | 50.28 | 51.2207 | 50.26 | 45.9 | 50 | 50.1732 | 48.6 | 50 | > 10 | 50.47 | 47 | 45.4 | 50 | 49.86 | 27
| 1 | 3 | 49.9 | 47.826 | 47.5 | 49.6 | 49.2 | 48.7 | 47.4 | 49.3 | 49.8 | 50.88 | 50.8292 | 49.36 | 45.2 | 51.5 | 50.8502 | 49.1 | 52 | > 10 | 48.74 | 48 | 45.4 | 50 | 51.28 | 16
| 1 | 4 | 49.6 | 49.049 | 46.1 | 48.5 | 48.6 | 50 | 47.6 | 49.6 | 51.5 | 49.31 | 51.9002 | 51.27 | 45.7 | 49.5 | 50.6221 | 50.2 | 52 | > 10 | 47.08 | 47 | 45.8 | 50 | 52.7 | 49.3686
| 1 | 5 | 49.9 | 48.72 | 48 | 48 | 47.6 | 49.6 | 47.4 | 48.9 | 50.5 | 50.46 | 51.6607 | 55.41 | 46.1 | 51 | 50.4066 | 49.6 | 50 | > 10 | 48.61 | 46 | 45.1 | 50 | 54.12 | 13
| 1 | 6 | 49.7 | 49.594 | 45.7 | 47.8 | 48.3 | 49 | 47 | 51.1 | 50.6 | 50.91 | 51.0939 | 54.95 | 45.4 | 50 | 50.5302 | 50.1 | 50 | > 10 | 47.9 | 47 | 44.5 | 49 | 54.54
| Mean |  | 49.4833 | 48.914 | 47.0167 | 48.2833 | 48.35 | 48.9667 | 47.75 | 50.1 | 50.5 | 50.5133 | 51.2556 | 51.68 | 45.75 | 50.1667 | 50.4287 | 49.5167 | 50.8333 |  | 48.7617 | 47 | 45.4667 | 50 | 51.8233
| Median |  | 49.65 | 49.024 | 47.35 | 48.2 | 48.35 | 49.25 | 47.5 | 49.9 | 50.5 | 50.67 | 51.1573 | 50.765 | 45.8 | 50 | 50.4684 | 49.55 | 50.5 |  | 48.675 | 47 | 45.4 | 50 | 51.99
| Std Dev. |  | 0.53072 | 0.60912 | 0.91086 | 0.76004 | 0.54314 | 1.0671 | 0.65651 | 1.0296 | 0.57619 | 0.68193 | 0.44065 | 2.8388 | 0.3937 | 0.93095 | 0.31094 | 0.6047 | 0.98319 |  | 1.2271 | 0.63246 | 0.70333 | 0.63246 | 2.4092
| Rel.Std.Dev. |  | 1.073% | 1.245% | 1.937% | 1.574% | 1.123% | 2.179% | 1.375% | 2.055% | 1.141% | 1.35% | 0.8597% | 5.493% | 0.8605% | 1.856% | 0.6166% | 1.221% | 1.934% |  | 2.517% | 1.346% | 1.547% | 1.265% | 4.649%
| PDM3 |  | 0.2324% | -0.9208% | -4.764% | -2.198% | -2.063% | -0.8142% | -3.279% | 1.481% | 2.292% | 2.319% | 3.822% | 4.682% | -7.33% | 1.617% | 2.147% | 0.2999% | 2.967% |  | -1.229% | -4.798% | -7.904% | 1.279% | 4.972%
| Z-Score (Absolute) |  | 0.01 | 0.37 | 1.55 | 0.76 | 0.72 | 0.33 | 1.09 | 0.37 | 0.62 | 0.63 | 1.09 | 1.36 | 2.34 | 0.42 | 0.58 | 0.01 | 0.83 | NA | 0.46 | 1.56 | 2.51 | 0.31 | 1.45
| Analytical results for Al in OREAS 610 (Certified Value 5.99 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 6.0348 | 5.81 | 6.11 | 5.91 | 5.91 | 5.63 | 5.89 | 5.98 | 6.13 | 6.2627 | 6.0119 | 4.7003 | 6.33 | 6.23 | 6.2165 | 6.29 | 5.247 | 5.03 | 6.07 | 5.63 | 5.74 | 3.1 | 5.817 | 1
| 1 | 2 | 5.8795 | 5.63 | 6.06 | 6.07 | 5.82 | 5.92 | 5.5 | 6.18 | 6.29 | 6.1129 | 6.1557 | 5.1081 | 6.4 | 6.22 | 6.2966 | 6.25 | 4.906 | 5.02 | 6.2 | 5.78 | 5.55 | 3.1 | 5.864 | #N/A
| 1 | 3 | 6.0476 | 5.68 | 5.66 | 6.02 | 5.87 | 5.87 | 5.57 | 5.91 | 6.22 | 6.1448 | 6.1771 | 5.0804 | 6.18 | 6.38 | 6.3437 | 6.09 | 5.223 | 5.24 | 6.1 | 5.86 | 5.72 | 3.1 | 5.911 | 16
| 1 | 4 | 5.9459 | 5.65 | 6.04 | 6.09 | 5.84 | 5.96 | 5.78 | 6 | 6.22 | 6.286 | 6.137 | 5.3295 | 6.28 | 6.29 | 6.3354 | 6.3 | 4.895 | 4.93 | 6.08 | 5.9 | 5.55 | 3.11 | 5.958 | 5.9878
| 1 | 5 | 6.094 | 5.66 | 6.65 | 5.99 | 5.83 | 5.95 | 5.52 | 5.95 | 6.1 | 6.2788 | 6.0171 | 5.9658 | 6.49 | 6.33 | 6.2895 | 6.27 | 4.895 | 5.01 | 6.15 | 5.63 | 5.9 | 3.1 | 6.005 | 14
| 1 | 6 | 5.9479 | 5.7 | 6.16 | 5.92 | 5.77 | 5.86 | 5.74 | 6.18 | 5.91 | 6.161 | 6.1416 | 5.2575 | 6.18 | 6.26 | 6.2518 | 6.16 | 5.155 | 5.16 | 5.98 | 5.73 | 5.98 | 3.1 | 6.052
| Mean |  | 5.9916 | 5.6883 | 6.1133 | 6 | 5.84 | 5.865 | 5.6667 | 6.0333 | 6.145 | 6.2077 | 6.1067 | 5.2403 | 6.31 | 6.285 | 6.2889 | 6.2267 | 5.0535 | 5.065 | 6.0967 | 5.755 | 5.74 | 3.1017 | 5.9345
| Median |  | 5.9913 | 5.67 | 6.085 | 6.005 | 5.835 | 5.895 | 5.655 | 5.99 | 6.175 | 6.2119 | 6.1393 | 5.1828 | 6.305 | 6.275 | 6.2931 | 6.26 | 5.0305 | 5.025 | 6.09 | 5.755 | 5.73 | 3.1 | 5.9345
| Std Dev. |  | 0.080066 | 0.064317 | 0.31747 | 0.074833 | 0.047329 | 0.12211 | 0.15921 | 0.11759 | 0.13398 | 0.076597 | 0.072832 | 0.41695 | 0.12296 | 0.061563 | 0.048659 | 0.083586 | 0.17232 | 0.11327 | 0.075011 | 0.11362 | 0.17652 | 0.0040825 | 0.087929
| Rel.Std.Dev. |  | 1.336% | 1.131% | 5.193% | 1.247% | 0.8104% | 2.082% | 2.81% | 1.949% | 2.18% | 1.234% | 1.193% | 7.957% | 1.949% | 0.9795% | 0.7737% | 1.342% | 3.41% | 2.236% | 1.23% | 1.974% | 3.075% | 0.1316% | 1.482%
| PDM3 |  | 0.06317% | -5.002% | 2.096% | 0.2032% | -2.469% | -2.051% | -5.364% | 0.7599% | 2.625% | 3.672% | 1.985% | -12.48% | 5.38% | 4.963% | 5.028% | 3.989% | -15.6% | -15.41% | 1.818% | -3.888% | -4.139% | -48.2% | -0.8907%
| Z-Score (Absolute) |  | 0 | 0.87 | 0.35 | 0.02 | 0.43 | 0.36 | 0.93 | 0.12 | 0.44 | 0.62 | 0.33 | 2.16 | 0.91 | 0.84 | 0.85 | 0.67 | 2.69 | 2.66 | 0.3 | 0.68 | 0.72 | 8.29 | 0.16
| Analytical results for As in OREAS 610 (Certified Value 2835 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 2818 | 2216.1 | 2690 | 2960 | 2900 | 2780 | 2900 | 3030 | 2710 | 2940.2 | 2961.264 | 2860.5 | 2940 | 3200 | 2881.9523 | 2890 | 2858.643 | 2624 | 2636 | 2780 | 966 | 2550 | 2394 | 1
| 1 | 2 | 2728 | 2215.9 | 2770 | 3040 | 2900 | 2940 | 2890 | 3120 | 2700 | 2871.1 | 2940.1869 | 2926.1 | 2850 | 3200 | 2891.5016 | 2800 | 2768.492 | 2672 | 2608 | 2686 | > 1000 | 2525 | 2406 | 17
| 1 | 3 | 2816 | 2120.7 | 2740 | 3010 | 2900 | 2950 | 2890 | 2980 | 2740 | 2910.5 | 3021.6234 | 2912.4 | 2920 | 3280 | 2894.8793 | 2836 | 3025.277 | 2702 | 2686 | 2759 | 835.5 | 2523 | 2418 | 16
| 1 | 4 | 2802 | 2238.4 | 2700 | 3050 | 2930 | 2990 | 2900 | 3020 | 2740 | 2835.4 | 2989.67 | 2882.4 | 2880 | 3180 | 2908.9964 | 2927 | 2817.554 | 2630 | 2573 | 2770 | > 1000 | 2584 | 2430 | 2834.5718
| 1 | 5 | 2835 | 2253.6 | 2780 | 3020 | 2880 | 2980 | 2840 | 2990 | 2730 | 2894.5 | 2894.6004 | 2969.2 | 2900 | 3320 | 2917.6468 | 2910 | 2867.244 | 2611 | 2547 | 2640 | > 1000 | 2534 | 2442 | 15
| 1 | 6 | 2797 | 2294 | 2720 | 2980 | 2880 | 2960 | 2890 | 3150 | 2770 | 2902.7 | 2949.3269 | 2958.2 | 2860 | 3250 | 2912.847 | 2948 | 2855.023 | 2684 | 2531 | 2680 | > 1000 | 2538 | 2454
| Mean |  | 2799.3333 | 2223.1167 | 2733.3333 | 3010 | 2898.3333 | 2933.3333 | 2885 | 3048.3333 | 2731.6667 | 2892.4 | 2959.4453 | 2918.1333 | 2891.6667 | 3238.3333 | 2901.3039 | 2885.1667 | 2865.3722 | 2653.8333 | 2596.8333 | 2719.1667 | 900.75 | 2542.3333 | 2424
| Median |  | 2809 | 2227.25 | 2730 | 3015 | 2900 | 2955 | 2890 | 3025 | 2735 | 2898.6 | 2955.2954 | 2919.25 | 2890 | 3225 | 2901.9379 | 2900 | 2856.833 | 2651 | 2590.5 | 2722.5 | 900.75 | 2536 | 2424
| Std Dev. |  | 37.4148 | 57.9271 | 36.697 | 34.641 | 18.3485 | 77.3736 | 22.5832 | 70.2614 | 24.8328 | 35.8487 | 43.4996 | 42.2089 | 34.8807 | 54.5588 | 13.9381 | 56.5842 | 86.4844 | 37.0212 | 58.3007 | 57.9186 | 92.2774 | 22.6156 | 22.4499
| Rel.Std.Dev. |  | 1.337% | 2.606% | 1.343% | 1.151% | 0.6331% | 2.638% | 0.7828% | 2.305% | 0.9091% | 1.239% | 1.47% | 1.446% | 1.206% | 1.685% | 0.4804% | 1.961% | 3.018% | 1.395% | 2.245% | 2.13% | 10.24% | 0.8896% | 0.9262%
| PDM3 |  | -1.243% | -21.57% | -3.572% | 6.189% | 2.249% | 3.484% | 1.779% | 7.541% | -3.63% | 2.04% | 4.405% | 2.948% | 2.014% | 14.24% | 2.354% | 1.785% | 1.087% | -6.376% | -8.387% | -4.071% | -68.22% | -10.31% | -14.48%
| Z-Score (Absolute) |  | 0.46 | 3.57 | 0.82 | 0.67 | 0.07 | 0.26 | 0 | 0.88 | 0.83 | 0.04 | 0.4 | 0.18 | 0.04 | 1.91 | 0.09 | 0 | 0.11 | 1.25 | 1.55 | 0.89 | 10.7 | 1.85 | 2.49
| Analytical results for Ba in OREAS 610 (Indicative Value 242 ppm) | [24 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 876 | 49 | 44 | 160 | 100 | 330 | 780 | 280 | 384 | 155.2838 | 48 | 1870 | 359.0475 | 2415 | 192.495 | 445 | 82 | 101 | > 1000 | 231 | 3373 |  |  | 1
| 1 | 2 | 696 | 29 | 40 | 160 | 70 | 980 | 320 | 240 | 372 | 146.7779 | 56 | 1930 | 459.5144 | 2578 | 233.653 | 458 | 64 | 135 | > 1000 | 234 | 3368 |  |  | 5
| 1 | 3 | 812 | 22 | 36 | 150 | 50 | 680 | 420 | 340 | 383 | 149.0632 | 176 | 1800 | 418.4477 | 2536 | 194.679 | 429 | 52 | 86 | > 1000 | 237 | 3363 |  |  | 16
| 1 | 4 | 1460 | 22 | 33 | 210 | 70 | 1030 | 500 | 360 | 395 | 152.5788 | 227 | 1900 | 435.8764 | 2566 | 193.045 | 431 | 62 | 119 | > 1000 | 231 | 3358 |  |  | 242.4673
| 1 | 5 | 1250 | 29 | 43 | 150 | 70 | 1090 | 290 | 360 | 391 | 152.7892 | 741 | 1920 | 403.2234 | 2579 | 208.253 | 436 | 55 | 120 | > 1000 | 230 | 3353 |  |  | 11
| 1 | 6 | 731 | 20 | 44 | 140 | 60 | 960 | 580 | 330 | 356 | 147.5135 | 71 | 2050 | 415.8347 | 2577 | 195.621 | 446 | 81 | 130 | > 1000 | 232 | 3348
| Mean |  | 970.8333 | 28.5 | 40 | 161.6667 | 70 | 845 | 481.6667 | 318.3333 | 380.1667 | 150.6677 | 219.8333 | 1911.6667 | 415.324 | 2541.8333 | 202.9577 | 440.8333 | 66 | 115.1667 |  | 232.5 | 3360.5
| Median |  | 844 | 25.5 | 41.5 | 155 | 70 | 970 | 460 | 335 | 383.5 | 150.821 | 123.5 | 1910 | 417.1412 | 2571.5 | 195.15 | 440.5 | 63 | 119.5 |  | 231.5 | 3360.5
| Std Dev. |  | 311.2853 | 10.7471 | 4.6043 | 24.8328 | 16.7332 | 289.1885 | 182.0348 | 48.3391 | 14.2185 | 3.3798 | 265.3702 | 82.3205 | 33.76 | 64.2352 | 16.1266 | 10.9438 | 12.7906 | 18.4544 |  | 2.5884 | 9.3541
| Rel.Std.Dev. |  | 32.06% | 37.71% | 11.51% | 15.36% | 23.9% | 34.22% | 37.79% | 15.19% | 3.74% | 2.243% | 120.7% | 4.306% | 8.129% | 2.527% | 7.946% | 2.483% | 19.38% | 16.02% |  | 1.113% | 0.2784%
| PDM3 |  | 300.4% | -88.25% | -83.5% | -33.32% | -71.13% | 248.5% | 98.65% | 31.29% | 56.79% | -37.86% | -9.335% | 688.4% | 71.29% | 948.3% | -16.29% | 81.81% | -72.78% | -52.5% |  | -4.111% | 1286%
| Z-Score (Absolute) |  | 2.53 | 0.9 | 0.86 | 0.41 | 0.75 | 2.07 | 0.75 | 0.16 | 0.38 | 0.45 | 0.2 | 5.95 | 0.51 | 8.24 | 0.26 | 0.6 | 0.76 | 0.58 | NA | 0.16 | 11.22
| Analytical results for Be in OREAS 610 (Certified Value 1.48 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.6 | 1 | 1.6 | 1.42 | 1.46 | 1.31 | 1.47 | 1.5 | 1.5 | 1.47 | 1.5798 | 2 | 1.88 | 1.5 | 1.2 | 1.267 | 1.4 | 1.5 | 1.4 | 2 | < 0.5 |  |  | 1
| 1 | 2 | 1.6 | 1 | 1.4 | 1.52 | 1.49 | 1.46 | 1.48 | 1.58 | 1.6 | 1.4 | 1.5709 | 2 | 1.9 | 1.4 | 1.2 | 1.308 | 1.5 | 1.5 | 1.3 | 2 | < 0.5 |  |  | 30
| 1 | 3 | 1.7 | 2 | 1.6 | 1.42 | 1.58 | 1.44 | 1.35 | 1.44 | 1.5 | 1.39 | 1.5288 | 2 | 1.98 | 1.4 | 1.2 | 1.391 | 1.6 | 1.5 | 1.4 | 2 | < 0.5 |  |  | 16
| 1 | 4 | 1.7 | 2 | 1.5 | 1.42 | 1.42 | 1.52 | 1.46 | 1.57 | 1.5 | 1.39 | 1.4769 | 2 | 1.94 | 1.8 | 1.1 | 1.25 | 1.5 | 1.5 | 1.4 | 2 | < 0.5 |  |  | 1.4751
| 1 | 5 | 1.7 | 2 | 1.6 | 1.64 | 1.53 | 1.46 | 1.39 | 1.42 | 1.5 | 1.4 | 1.4889 | 2 | 1.97 | 1.5 | 1 | 1.285 | 1.5 | 1.5 | 1.3 | 2 | < 0.5 |  |  | 16
| 1 | 6 | 1.6 | 1 | 1.6 | 1.29 | 1.43 | 1.52 | 1.39 | 1.55 | 1.5 | 1.43 | 1.431 | 2 | 1.93 | 1.7 | 1.3 | 1.279 | 1.5 | 1.3 | 1.3 | 2 | < 0.5
| Mean |  | 1.65 | 1.5 | 1.55 | 1.4517 | 1.485 | 1.4517 | 1.4233 | 1.51 | 1.5167 | 1.4133 | 1.5127 | 2 | 1.9333 | 1.55 | 1.1667 | 1.2967 | 1.5 | 1.4667 | 1.35 | 2
| Median |  | 1.65 | 1.5 | 1.6 | 1.42 | 1.475 | 1.46 | 1.425 | 1.525 | 1.5 | 1.4 | 1.5089 | 2 | 1.935 | 1.5 | 1.2 | 1.282 | 1.5 | 1.5 | 1.35 | 2
| Std Dev. |  | 0.054772 | 0.54772 | 0.083666 | 0.11771 | 0.061563 | 0.07705 | 0.053541 | 0.068118 | 0.040825 | 0.031411 | 0.057736 | 0 | 0.038816 | 0.16432 | 0.10328 | 0.050067 | 0.063246 | 0.08165 | 0.054772 | 0
| Rel.Std.Dev. |  | 3.32% | 36.51% | 5.398% | 8.109% | 4.146% | 5.308% | 3.762% | 4.511% | 2.692% | 2.222% | 3.817% | 0% | 2.008% | 10.6% | 8.853% | 3.861% | 4.216% | 5.567% | 4.057% | 0%
| PDM3 |  | 11.86% | 1.691% | 5.081% | -1.586% | 0.674% | -1.586% | -3.507% | 2.369% | 2.821% | -4.185% | 2.554% | 35.59% | 31.07% | 5.081% | -20.91% | -12.09% | 1.691% | -0.5689% | -8.478% | 35.59%
| Z-Score (Absolute) |  | 2.04 | NA | 0.87 | 0.28 | 0.11 | 0.28 | 0.62 | 0.4 | 0.48 | 0.73 | 0.43 | NA | 5.37 | 0.87 | 3.63 | 2.1 | 0.28 | 0.11 | 1.48 | NA | 14.38
Results from laboratories 3, 14 & 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 610 (Certified Value 224 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 207.25 | 218.1 | 192 | 221 | 227 | 212 | 222 | 226 | 250 | 232.81 | 235.5299 | 241.04 | 222 | 222 | 224.0767 | 225.28 | 226.345 | 216.2 | 215 | 207 | 235 | 192 | 146 | 1
| 1 | 2 | 211.21 | 220.46 | 193 | 225 | 222 | 222 | 220 | 237 | 246 | 229.07 | 236.2944 | 247.84 | 221 | 229 | 218.8552 | 212.36 | 214.91 | 218.5 | 219 | 202 | 230.3 | 191 | 145.4 | 31
| 1 | 3 | 210.96 | 222.61 | 189 | 229 | 229 | 225 | 218 | 225 | 248 | 226.18 | 241.2563 | 247 | 217 | 234 | 217.1335 | 218.68 | 226.721 | 215 | 221 | 207 | 231.5 | 190 | 144.8 | 16
| 1 | 4 | 209.39 | 223.19 | 190 | 230 | 228 | 225 | 217 | 221 | 255 | 227.72 | 239.0416 | 242.2 | 219 | 231 | 214.8602 | 229.47 | 219.101 | 217.9 | 212 | 211 | 228.1 | 190 | 144.2 | 223.5764
| 1 | 5 | 208.58 | 236.33 | 199 | 225 | 222 | 227 | 220 | 224 | 257 | 229.01 | 239.5097 | 239.7 | 216 | 236 | 221.3604 | 223.22 | 220.715 | 217.6 | 212 | 200 | 227.7 | 191 | 146.6 | 17
| 1 | 6 | 207.33 | 237.14 | 193 | 226 | 223 | 224 | 218 | 240 | 252 | 230.07 | 239.502 | 236.42 | 216 | 232 | 222.7719 | 213.16 | 224.243 | 219.3 | 213 | 204 | 224 | 188 | 146
| Mean |  | 209.12 | 226.305 | 192.6667 | 226 | 225.1667 | 222.5 | 219.1667 | 228.8333 | 251.3333 | 229.1433 | 238.5223 | 242.3667 | 218.5 | 230.6667 | 219.843 | 220.3617 | 222.0058 | 217.4167 | 215.3333 | 205.1667 | 229.4333 | 190.3333 | 145.5
| Median |  | 208.985 | 222.9 | 192.5 | 225.5 | 225 | 224.5 | 219 | 225.5 | 251 | 229.04 | 239.2718 | 241.62 | 218 | 231.5 | 220.1078 | 220.95 | 222.479 | 217.75 | 214 | 205.5 | 229.2 | 190.5 | 145.7
| Std Dev. |  | 1.7216 | 8.2794 | 3.5024 | 3.2249 | 3.1885 | 5.3944 | 1.8348 | 7.7309 | 4.1793 | 2.242 | 2.1724 | 4.3746 | 2.5884 | 4.8854 | 3.52 | 6.8416 | 4.6157 | 1.569 | 3.8297 | 3.9707 | 3.75 | 1.3663 | 0.88318
| Rel.Std.Dev. |  | 0.8233% | 3.659% | 1.818% | 1.427% | 1.416% | 2.424% | 0.8372% | 3.378% | 1.663% | 0.9784% | 0.9108% | 1.805% | 1.185% | 2.118% | 1.601% | 3.105% | 2.079% | 0.7216% | 1.779% | 1.935% | 1.634% | 0.7178% | 0.607%
| PDM3 |  | -6.466% | 1.22% | -13.83% | 1.084% | 0.7113% | -0.4814% | -1.972% | 2.351% | 12.41% | 2.49% | 6.685% | 8.404% | -2.271% | 3.171% | -1.67% | -1.438% | -0.7025% | -2.755% | -3.687% | -8.234% | 2.62% | -14.87% | -34.92%
| Z-Score (Absolute) |  | 1.27 | 0.42 | 2.9 | 0.39 | 0.31 | 0.05 | 0.28 | 0.67 | 2.9 | 0.7 | 1.63 | 2.01 | 0.35 | 0.86 | 0.21 | 0.16 | 0 | 0.45 | 0.66 | 1.66 | 0.73 | 3.13 | 7.56
| Analytical results for Ca in OREAS 610 (Certified Value 0.241 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.2411 | 0.23 | 0.24 | 0.24 | 0.25 | 0.23 | 0.24 | 0.25 | 0.241 | 0.2575 | 0.26145 | 0.1702 | 0.26 | 0.24 | 0.24389 | 0.19 | 0.192 | 0.21 | 0.24 | 0.21 | 0.24 | 0.14 | 0.2747 | 1
| 1 | 2 | 0.2354 | 0.22 | 0.23 | 0.25 | 0.25 | 0.25 | 0.23 | 0.25 | 0.251 | 0.2507 | 0.25601 | 0.1964 | 0.26 | 0.24 | 0.24888 | 0.22 | 0.19 | 0.22 | 0.24 | 0.22 | 0.24 | 0.14 | 0.2675 | #N/A
| 1 | 3 | 0.2424 | 0.22 | 0.23 | 0.24 | 0.25 | 0.25 | 0.23 | 0.25 | 0.244 | 0.2507 | 0.25711 | 0.1874 | 0.25 | 0.25 | 0.25124 | 0.22 | 0.191 | 0.23 | 0.25 | 0.23 | 0.24 | 0.14 | 0.2603 | 16
| 1 | 4 | 0.2387 | 0.22 | 0.23 | 0.25 | 0.25 | 0.25 | 0.23 | 0.25 | 0.246 | 0.2574 | 0.26012 | 0.1964 | 0.25 | 0.25 | 0.24896 | 0.2 | 0.191 | 0.22 | 0.24 | 0.22 | 0.23 | 0.14 | 0.2531 | 0.24074
| 1 | 5 | 0.2441 | 0.22 | 0.23 | 0.24 | 0.25 | 0.25 | 0.22 | 0.25 | 0.243 | 0.2561 | 0.25211 | 0.2268 | 0.26 | 0.25 | 0.24919 | 0.23 | 0.192 | 0.21 | 0.25 | 0.21 | 0.24 | 0.14 | 0.2459 | 18
| 1 | 6 | 0.2391 | 0.22 | 0.23 | 0.24 | 0.25 | 0.25 | 0.23 | 0.26 | 0.241 | 0.2509 | 0.25435 | 0.1993 | 0.25 | 0.25 | 0.24528 | 0.23 | 0.195 | 0.22 | 0.24 | 0.21 | 0.25 | 0.14 | 0.2387
| Mean |  | 0.24013 | 0.22167 | 0.23167 | 0.24333 | 0.25 | 0.24667 | 0.23 | 0.25167 | 0.24433 | 0.25388 | 0.25686 | 0.19608 | 0.255 | 0.24667 | 0.24791 | 0.215 | 0.19183 | 0.21833 | 0.24333 | 0.21667 | 0.24 | 0.14 | 0.2567
| Median |  | 0.2401 | 0.22 | 0.23 | 0.24 | 0.25 | 0.25 | 0.23 | 0.25 | 0.2435 | 0.2535 | 0.25656 | 0.1964 | 0.255 | 0.25 | 0.24892 | 0.22 | 0.1915 | 0.22 | 0.24 | 0.215 | 0.24 | 0.14 | 0.2567
| Std Dev. |  | 0.0030768 | 0.0040825 | 0.0040825 | 0.005164 | 0 | 0.008165 | 0.0063246 | 0.0040825 | 0.0037771 | 0.0034505 | 0.0035031 | 0.018436 | 0.0054772 | 0.005164 | 0.0027518 | 0.016432 | 0.0017224 | 0.0075277 | 0.005164 | 0.008165 | 0.0063246 | 0 | 0.01347
| Rel.Std.Dev. |  | 1.281% | 1.842% | 1.762% | 2.122% | 0% | 3.31% | 2.75% | 1.622% | 1.546% | 1.359% | 1.364% | 9.402% | 2.148% | 2.094% | 1.11% | 7.643% | 0.8979% | 3.448% | 2.122% | 3.768% | 2.635% | 0% | 5.247%
| PDM3 |  | -0.2524% | -7.923% | -3.769% | 1.077% | 3.846% | 2.461% | -4.462% | 4.538% | 1.492% | 5.459% | 6.695% | -18.55% | 5.923% | 2.461% | 2.977% | -10.69% | -20.32% | -9.308% | 1.077% | -10% | -0.3077% | -41.85% | 6.629%
| Z-Score (Absolute) |  | 0.18 | 1.25 | 0.67 | 0 | 0.39 | 0.19 | 0.77 | 0.48 | 0.06 | 0.61 | 0.78 | 2.73 | 0.67 | 0.19 | 0.26 | 1.64 | 2.98 | 1.44 | 0 | 1.54 | 0.19 | 5.97 | 0.77
| Analytical results for Cd in OREAS 610 (Certified Value 12.1 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 11.35 | 13.09 | 11.2 | 11.5 | 12 | 11.5 | 12.3 | 12.6 | 11.5 | 12.98 | 12.0025 | 12.34 | 12.6 | 12.5 | 12.5007 | 10.2 | 12.268 | 12.4 | 10.62 | 12 | 11.05 | 28 | 13.76 | 1
| 1 | 2 | 11 | 12.52 | 11.6 | 11.9 | 12.1 | 12.2 | 12.1 | 13 | 11.8 | 12.57 | 11.9712 | 12.2 | 12.4 | 12.5 | 12.3127 | 9.85 | 12.022 | 12.5 | 10.68 | 11 | 10.74 | 29 | 14.1 | 32
| 1 | 3 | 11.41 | 12.97 | 11.6 | 12.5 | 12.4 | 11.9 | 12.05 | 12.25 | 11.6 | 12.61 | 11.8124 | 12.44 | 12.5 | 12.5 | 12.4292 | 10.14 | 12.891 | 12.7 | 11 | 12 | 10.73 | 29 | 14.44 | 16
| 1 | 4 | 11.4 | 14.07 | 11.4 | 11.95 | 12.35 | 12.05 | 12.45 | 12.6 | 11.7 | 12.2 | 12.026 | 12.45 | 13.2 | 13 | 12.4843 | 10.42 | 12.233 | 12.8 | 10.89 | 11 | 10.6 | 29 | 14.78 | 12.0867
| 1 | 5 | 11.46 | 12.84 | 12.4 | 12.05 | 11.9 | 12.25 | 11.85 | 12.35 | 11.9 | 12.21 | 11.8941 | 13.88 | 12.6 | 13 | 12.6225 | 10.32 | 12.248 | 12.7 | 10.66 | 11 | 10.59 | 28 | 15.12 | 19
| 1 | 6 | 11.18 | 13.67 | 11.7 | 11.95 | 12.15 | 11.85 | 11.9 | 12.9 | 12 | 12.41 | 11.8551 | 14.59 | 12.3 | 13 | 12.5085 | 10.03 | 12.3 | 12.8 | 10.52 | 11 | 10.2 | 29 | 15.06
| Mean |  | 11.3 | 13.1933 | 11.65 | 11.975 | 12.15 | 11.9583 | 12.1083 | 12.6167 | 11.75 | 12.4967 | 11.9269 | 12.9833 | 12.6 | 12.75 | 12.4763 | 10.16 | 12.327 | 12.65 | 10.7283 | 11.3333 | 10.6517 | 28.6667 | 14.5433
| Median |  | 11.375 | 13.03 | 11.6 | 11.95 | 12.125 | 11.975 | 12.075 | 12.6 | 11.75 | 12.49 | 11.9326 | 12.445 | 12.55 | 12.75 | 12.4925 | 10.17 | 12.258 | 12.7 | 10.67 | 11 | 10.665 | 29 | 14.61
| Std Dev. |  | 0.17584 | 0.57176 | 0.40866 | 0.32055 | 0.19494 | 0.27462 | 0.23112 | 0.29439 | 0.18708 | 0.2931 | 0.085846 | 0.99925 | 0.31623 | 0.27386 | 0.10205 | 0.20406 | 0.2934 | 0.16432 | 0.18005 | 0.5164 | 0.2768 | 0.5164 | 0.54338
| Rel.Std.Dev. |  | 1.556% | 4.334% | 3.508% | 2.677% | 1.604% | 2.296% | 1.909% | 2.333% | 1.592% | 2.345% | 0.7198% | 7.696% | 2.51% | 2.148% | 0.8179% | 2.008% | 2.38% | 1.299% | 1.678% | 4.556% | 2.599% | 1.801% | 3.736%
| PDM3 |  | -6.509% | 9.156% | -3.613% | -0.9243% | 0.5235% | -1.062% | 0.1788% | 4.385% | -2.786% | 3.392% | -1.323% | 7.418% | 4.247% | 5.488% | 3.223% | -15.94% | 1.988% | 4.66% | -11.24% | -6.233% | -11.87% | 137.2% | 20.32%
| Z-Score (Absolute) |  | 1.23 | 1.51 | 0.72 | 0.25 | 0 | 0.28 | 0.06 | 0.67 | 0.58 | 0.5 | 0.32 | 1.2 | 0.65 | 0.87 | 0.47 | 2.88 | 0.26 | 0.72 | 2.05 | NA | 2.17 | NA | 3.46
Results from laboratories 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 610 (Certified Value 46.7 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 1
2
| 1 | 1 | 37.62 | 36.4 | 43.8 | 43.6 | 51.3 | 47.6 | 50.3 | 49.81 | 27.76 | 51.9 | 54.9 | 47.059 | 45.1 | 36.9 | 48.5 | 49 | [7 empty] | 1
| 1 | 2 | 38.19 | 30.3 | 44.5 | 44.8 | 48.4 | 47 | 51.4 | 49.53 | 29.79 | 50.6 | 51.2 | 45.4153 | 43.9 | 38.01 | 49.02 | 48 | [7 empty] | 33
| 1 | 3 | 39.38 | 29.6 | 46.9 | 48.5 | 47.9 | 47.2 | 48.9 | 50.04 | 27.83 | 51 | 53.3 | 45.4048 | 40.7 | 39.06 | 48.22 | 48 | [7 empty] | 16
| 1 | 4 | 38.71 | 33 | 44.3 | 45.8 | 49.5 | 48.1 | 50.1 | 50.12 | 29.39 | 48.3 | 53.2 | 44.972 | 44.3 | 37.12 | 49.04 | 47 | [7 empty] | 46.6884
| 1 | 5 | 38.59 | 35.3 | 46.1 | 44.1 | 49.8 | 45.7 | 50.9 | 49.48 | 31.39 | 49.1 | 52.2 | 45.7902 | 46.2 | 36.13 | 45.8 | 48 | [7 empty] | 20
| 1 | 6 | 38.2 | 36 | 48 | 45.7 | 47 | 45 | 50.9 | 51.19 | 30.06 | 48.6 | 54.3 | 47.8223 | 44.4 | 35.83 | 49.45 | 48
| Mean |  | 38.4483 | 33.4333 | 45.6 | 45.4167 | 48.9833 | 46.7667 | 50.4167 | 50.0283 | 29.37 | 49.9167 | 53.1833 | 46.0773 | 44.1 | 37.175 | 48.3383 | 48
| Median |  | 38.395 | 34.15 | 45.3 | 45.25 | 48.95 | 47.1 | 50.6 | 49.925 | 29.59 | 49.85 | 53.25 | 45.6028 | 44.35 | 37.01 | 48.76 | 48
| Std Dev. |  | 0.59516 | 2.9521 | 1.6637 | 1.7406 | 1.5329 | 1.181 | 0.87731 | 0.62525 | 1.3926 | 1.4552 | 1.3497 | 1.1139 | 1.8515 | 1.2025 | 1.3173 | 0.63246
| Rel.Std.Dev. |  | 1.548% | 8.83% | 3.649% | 3.833% | 3.129% | 2.525% | 1.74% | 1.25% | 4.742% | 2.915% | 2.538% | 2.417% | 4.198% | 3.235% | 2.725% | 1.318%
| PDM3 |  | -17.65% | -28.39% | -2.331% | -2.724% | 4.915% | 0.1675% | 7.985% | 7.154% | -37.09% | 6.914% | 13.91% | -1.309% | -5.544% | -20.38% | 3.534% | 2.809%
| Z-Score (Absolute) |  | 1.78 | 2.89 | 0.18 | 0.22 | 0.57 | 0.08 | 0.89 | 0.8 | 3.8 | 0.78 | 1.51 | 0.08 | 0.52 | 2.06 | 0.43 | 0.35
| Analytical results for Co in OREAS 610 (Certified Value 7.72 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 8 | 8.3 | 7.4 | 7.7 | 7.6 | 7.1 | 7.9 | 7.2 | 8 | 8.1 | 7.6335 | 7.5 | 8.13 | 10 | 7.3486 | 7.95 | 7.646 | 7.6 | 7.7 | 7 | 8 | 7 | 4.0386 | 1
| 1 | 2 | 8 | 8.1 | 7.6 | 7.8 | 7.8 | 7.8 | 7.8 | 7.6 | 8 | 7.9 | 7.8492 | 7.3 | 8.15 | 10 | 7.2749 | 7.56 | 7.491 | 7.7 | 7.5 | 7 | 7.8 | 8 | 4.569 | 21
| 1 | 3 | 8.1 | 7.5 | 7.7 | 8.2 | 7.7 | 7.4 | 7.9 | 7.2 | 7 | 8 | 7.7386 | 7.6 | 8.22 | 10 | 7.1317 | 7.66 | 8.47 | 7.9 | 7.9 | 7 | 7.8 | 7 | 4.52 | 16
| 1 | 4 | 8.3 | 7.7 | 7.3 | 7.4 | 8 | 7.5 | 7.6 | 7.4 | 8 | 7.9 | 7.3749 | 7.4 | 8.15 | 15 | 7.6992 | 8.05 | 7.985 | 7.6 | 7.6 | 7 | 8.1 | 7 | 4.535 | 7.7238
| 1 | 5 | 8.4 | 7.4 | 7.9 | 7.6 | 7.6 | 7.5 | 7.5 | 7 | 8 | 7.9 | 7.4245 | 7.6 | 7.83 | 15 | 7.2926 | 7.87 | 8.275 | 7.4 | 7.4 | 7 | 7.9 | 8 | 5.082 | 21
| 1 | 6 | 8 | 8 | 7.8 | 7.9 | 7.8 | 7.3 | 7.3 | 7.4 | 8 | 7.9 | 7.6374 | 7.9 | 8.04 | 15 | 7.4184 | 7.75 | 7.708 | 7.6 | 7.6 | 7 | 8 | 8 | 4.699
| Mean |  | 8.1333 | 7.8333 | 7.6167 | 7.7667 | 7.75 | 7.4333 | 7.6667 | 7.3 | 7.8333 | 7.95 | 7.6097 | 7.55 | 8.0867 | 12.5 | 7.3609 | 7.8067 | 7.9292 | 7.6333 | 7.6167 | 7 | 7.9333 | 7.5 | 4.5739
| Median |  | 8.05 | 7.85 | 7.65 | 7.75 | 7.75 | 7.45 | 7.7 | 7.3 | 8 | 7.9 | 7.6355 | 7.55 | 8.14 | 12.5 | 7.3206 | 7.81 | 7.8465 | 7.6 | 7.6 | 7 | 7.95 | 7.5 | 4.552
| Std Dev. |  | 0.17512 | 0.3559 | 0.23166 | 0.27325 | 0.15166 | 0.23381 | 0.24221 | 0.20976 | 0.40825 | 0.083666 | 0.18152 | 0.20736 | 0.13837 | 2.7386 | 0.19105 | 0.18403 | 0.38375 | 0.1633 | 0.17224 | 0 | 0.12111 | 0.54772 | 0.3361
| Rel.Std.Dev. |  | 2.153% | 4.543% | 3.041% | 3.518% | 1.957% | 3.145% | 3.159% | 2.873% | 5.212% | 1.052% | 2.385% | 2.747% | 1.711% | 21.91% | 2.595% | 2.357% | 4.84% | 2.139% | 2.261% | 0% | 1.527% | 7.303% | 7.348%
| PDM3 |  | 5.303% | 1.419% | -1.387% | 0.5555% | 0.3397% | -3.76% | -0.7392% | -5.486% | 1.419% | 2.929% | -1.477% | -2.25% | 4.699% | 61.84% | -4.698% | 1.073% | 2.659% | -1.171% | -1.387% | -9.371% | 2.713% | -2.897% | -40.78%
| Z-Score (Absolute) |  | 1.89 | 0.67 | 0.2 | 0.4 | 0.34 | 0.94 | 0 | 1.48 | NA | 1.15 | 0.23 | 0.47 | 1.7 | NA | 1.24 | 0.57 | 1.06 | 0.13 | 0.2 | NA | 1.08 | NA | 12.51
Results from laboratories 10, 16, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 610 (Certified Value 39.1 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
1
| 1 | 1 | 35 | 30 | 30 | 41 | 40 | 40 | 41 | 41 | 39 | 46 | 41.0445 | 51 | 36 | 20 | 36.0727 | 41 | 30.209 | 39 | 33 | 31 | 46 | 44 | 45.42 | 1
| 1 | 2 | 39 | 34 | 26 | 40 | 40 | 37 | 40 | 42 | 40 | 45 | 39.7158 | 41 | 35.9 | 30 | 41.5669 | 39 | 30.943 | 39 | 38 | 37 | 48 | 44 | 45.22 | 35
| 1 | 3 | 39 | 33 | 28 | 40 | 40 | 38 | 39 | 40 | 41 | 46 | 39.2581 | 40 | 36.2 | 30 | 40.0842 | 40 | 32.095 | 40 | 30 | 43 | 48 | 43 | 45.02 | 16
| 1 | 4 | 37 | 34 | 29 | 39 | 40 | 39 | 41 | 41 | 41 | 47 | 39.601 | 45 | 38 | 20 | 31.113 | 41 | 29.537 | 40 | 33 | 35 | 46 | 43 | 44.82 | 39.1331
| 1 | 5 | 39 | 32 | 42 | 39 | 40 | 38 | 40 | 42 | 39 | 47 | 40.3905 | 50 | 39.3 | 20 | 24.5082 | 41 | 34.258 | 39 | 31 | 38 | 39 | 44 | 44.62 | 22
| 1 | 6 | 37 | 35 | 42 | 38 | 40 | 38 | 40 | 43 | 40 | 46 | 41.1503 | 50 | 45.5 | 30 | 33.751 | 40 | 31.27 | 40 | 30 | 35 | 44 | 43 | 44.42
| Mean |  | 37.6667 | 33 | 32.8333 | 39.5 | 40 | 38.3333 | 40.1667 | 41.5 | 40 | 46.1667 | 40.1934 | 46.1667 | 38.4833 | 25 | 34.516 | 40.3333 | 31.3853 | 39.5 | 32.5 | 36.5 | 45.1667 | 43.5 | 44.92
| Median |  | 38 | 33.5 | 29.5 | 39.5 | 40 | 38 | 40 | 41.5 | 40 | 46 | 40.0532 | 47.5 | 37.1 | 25 | 34.9118 | 40.5 | 31.1065 | 39.5 | 32 | 36 | 46 | 43.5 | 44.92
| Std Dev. |  | 1.633 | 1.7889 | 7.2226 | 1.0488 | 0 | 1.0328 | 0.75277 | 1.0488 | 0.89443 | 0.75277 | 0.79154 | 4.8751 | 3.6929 | 5.4772 | 6.2524 | 0.8165 | 1.6592 | 0.54772 | 3.0166 | 3.9875 | 3.3714 | 0.54772 | 0.37417
| Rel.Std.Dev. |  | 4.335% | 5.421% | 22% | 2.655% | 0% | 2.694% | 1.874% | 2.527% | 2.236% | 1.631% | 1.969% | 10.56% | 9.596% | 21.91% | 18.11% | 2.024% | 5.287% | 1.387% | 9.282% | 10.92% | 7.464% | 1.259% | 0.833%
| PDM3 |  | -3.747% | -15.67% | -16.1% | 0.9376% | 2.215% | -2.044% | 2.641% | 6.048% | 2.215% | 17.97% | 2.709% | 17.97% | -1.66% | -36.12% | -11.8% | 3.067% | -19.8% | 0.9376% | -16.95% | -6.729% | 15.42% | 11.16% | 14.79%
| Z-Score (Absolute) |  | 0.53 | 1.71 | 1.75 | 0.06 | 0.06 | 0.36 | 0.11 | 0.44 | 0.06 | 1.62 | 0.11 | 1.62 | 0.32 | NA | 1.32 | 0.15 | 2.12 | 0.06 | 1.83 | 0.82 | 1.37 | 0.95 | 1.31
Results from laboratory 16 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 610 (Certified Value 2.16 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 2 | 2.2 | 2.24 | 2.21 | 2.03 | 2.09 | 2.22 | 2.19 | 2.1 | 2.26 | 2.451 | 2 | 2.02 | 2.2 | 1.87 | 1.87 | 2 | 2.5 | [5 empty] | 1
| 1 | 2 | 2 | 2.1 | 2.05 | 2.26 | 2.04 | 2.2 | 2.25 | 2.27 | 2.2 | 2.22 | 2.3654 | 2.1 | 2.02 | 2.2 | 1.95 | 2.03 | 2 | 2.4 | [5 empty] | 36
| 1 | 3 | 2 | 2.2 | 2.09 | 2.32 | 2.09 | 2.12 | 2.13 | 2.12 | 2.2 | 2.3 | 2.4076 | 2 | 2.04 | 2.2 | 1.96 | 2.02 | 2 | 2.4 | [5 empty] | 16
| 1 | 4 | 1.9 | 2.2 | 2.04 | 2.21 | 2.06 | 2.16 | 2.3 | 2.19 | 2.2 | 2.3 | 2.3285 | 2.1 | 2.03 | 2.3 | 1.92 | 1.94 | 2 | 2.3 | [5 empty] | 2.1551
| 1 | 5 | 2 | 2.3 | 2.23 | 2.31 | 2.06 | 2.18 | 2.12 | 2.11 | 2.3 | 2.29 | 2.2769 | 2.3 | 1.97 | 2.3 | 1.98 | 2.01 | 2 | 2.3 | [5 empty] | 23
| 1 | 6 | 1.9 | 2.1 | 2.18 | 2.23 | 1.99 | 2.1 | 2.14 | 2.21 | 2.4 | 2.28 | 2.4305 | 2.2 | 2.02 | 2.2 | 1.93 | 1.89 | 2 | 2.4
| Mean |  | 1.9667 | 2.1833 | 2.1383 | 2.2567 | 2.045 | 2.1417 | 2.1933 | 2.1817 | 2.2333 | 2.275 | 2.3766 | 2.1167 | 2.0167 | 2.2333 | 1.935 | 1.96 | 2 | 2.3833
| Median |  | 2 | 2.2 | 2.135 | 2.245 | 2.05 | 2.14 | 2.18 | 2.19 | 2.2 | 2.285 | 2.3865 | 2.1 | 2.02 | 2.2 | 1.94 | 1.975 | 2 | 2.4
| Std Dev. |  | 0.05164 | 0.075277 | 0.089759 | 0.048854 | 0.033912 | 0.044907 | 0.074207 | 0.05947 | 0.10328 | 0.030822 | 0.065992 | 0.1169 | 0.024221 | 0.05164 | 0.038341 | 0.069857 | 0 | 0.075277
| Rel.Std.Dev. |  | 2.626% | 3.448% | 4.198% | 2.165% | 1.658% | 2.097% | 3.383% | 2.726% | 4.624% | 1.355% | 2.777% | 5.523% | 1.201% | 2.312% | 1.981% | 3.564% | 0% | 3.158%
| PDM3 |  | -8.743% | 1.31% | -0.7778% | 4.713% | -5.109% | -0.6232% | 1.774% | 1.233% | 3.63% | 5.564% | 10.28% | -1.783% | -6.423% | 3.63% | -10.21% | -9.053% | -7.197% | 10.59%
| Z-Score (Absolute) |  | 1.93 | 0.01 | 0.39 | 0.67 | 1.23 | 0.36 | 0.1 | 0 | 0.46 | 0.84 | 1.75 | 0.58 | 1.48 | 0.46 | 2.22 | 1.99 | NA | 1.81
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 610 (Certified Value 0.971 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*AAS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 3
3
| 1 | 1 | 0.9591 | 0.94863 | 0.969 | 0.938 | 0.971 | 0.916 | 0.984 | 0.968 | 0.992 | 1.0163 | 0.97601 | 0.97136 | 1.01 | 0.963 | 0.97591 | 0.9639 | 1 | 0.9707 | 0.9474 | 0.9128 | 0.97608 | 1.1 | 1.0474 | 1
| 1 | 2 | 0.9346 | 0.95133 | 0.971 | 0.933 | 0.971 | 0.956 | 0.965 | 0.999 | 0.994 | 0.98974 | 0.96687 | 0.97136 | 1.01 | 0.97 | 0.97687 | 0.9606 | 1 | 0.9715 | 0.9714 | 0.9352 | 0.98133 | 1.22 | 1.047 | 23
| 1 | 3 | 0.968 | 0.94666 | 1.02 | 0.946 | 0.975 | 0.936 | 0.969 | 0.941 | 1.01 | 0.98938 | 0.9915 | 0.97729 | 0.999 | 0.997 | 0.97902 | 0.9635 | 1.01 | 0.9723 | 0.9021 | 0.9511 | 0.9694 | 1.05 | 1.0366 | 16
| 1 | 4 | 0.9481 | 0.95667 | 0.969 | 0.945 | 0.966 | 0.959 | 0.985 | 0.952 | 1.02 | 0.98708 | 0.98516 | 0.98657 | 0.997 | 0.959 | 0.99092 | 0.9636 | 1 | 0.9741 | 0.9551 | 0.9548 | 0.98377 | 1.08 | 1.0562 | 0.97124
| 1 | 5 | 0.9826 | 0.963 | 0.948 | 0.942 | 0.956 | 0.956 | 0.969 | 0.958 | 1.07 | 1.0124 | 0.97773 | 0.97453 | 0.997 | 0.989 | 0.98502 | 0.9637 | 1.01 | 0.9797 | 0.9457 | 0.9292 | 0.99437 | 1.1 | 1.0558 | 24
| 1 | 6 | 0.9354 | 0.96268 | 0.979 | 0.927 | 0.961 | 0.941 | 0.969 | 0.993 | 0.996 | 1.013 | 0.96886 | 0.99047 | 1.01 | 0.957 | 0.97712 | 0.9614 | 1.01 | 0.9767 | 0.9451 | 0.9263 | 0.98284 | 1.11 | 1.0954
| Mean |  | 0.95463 | 0.95483 | 0.976 | 0.9385 | 0.96667 | 0.944 | 0.9735 | 0.9685 | 1.0137 | 1.0013 | 0.97769 | 0.9786 | 1.0038 | 0.9725 | 0.98081 | 0.96278 | 1.005 | 0.97417 | 0.94447 | 0.9349 | 0.9813 | 1.11 | 1.0564
| Median |  | 0.9536 | 0.954 | 0.97 | 0.94 | 0.9685 | 0.9485 | 0.969 | 0.963 | 1.003 | 1.0011 | 0.97687 | 0.97591 | 1.0045 | 0.9665 | 0.97807 | 0.96355 | 1.005 | 0.9732 | 0.94655 | 0.9322 | 0.98208 | 1.1 | 1.0516
| Std Dev. |  | 0.018946 | 0.0070597 | 0.023883 | 0.0073959 | 0.0071181 | 0.016553 | 0.008666 | 0.023106 | 0.029649 | 0.013891 | 0.009428 | 0.0080936 | 0.0067946 | 0.016682 | 0.0059347 | 0.0014106 | 0.0054772 | 0.0034541 | 0.022995 | 0.015833 | 0.0083445 | 0.057966 | 0.02041
| Rel.Std.Dev. |  | 1.985% | 0.7394% | 2.447% | 0.7881% | 0.7364% | 1.753% | 0.8902% | 2.386% | 2.925% | 1.387% | 0.9643% | 0.8271% | 0.6769% | 1.715% | 0.6051% | 0.1465% | 0.545% | 0.3546% | 2.435% | 1.694% | 0.8504% | 5.222% | 1.932%
| PDM3 |  | -1.71% | -1.69% | 0.4899% | -3.371% | -0.4711% | -2.805% | 0.2325% | -0.2823% | 4.368% | 3.098% | 0.6639% | 0.7573% | 3.356% | 0.1295% | 0.9853% | -0.8709% | 3.476% | 0.3012% | -2.757% | -3.742% | 1.035% | 14.29% | 8.768%
| Z-Score (Absolute) |  | 0.68 | 0.67 | 0.06 | 1.24 | 0.26 | 1.05 | 0.02 | 0.2 | 1.38 | 0.95 | 0.12 | 0.15 | 1.03 | 0.06 | 0.23 | 0.4 | 1.08 | 0 | 1.04 | 1.37 | 0.25 | 4.74 | 2.87
| Analytical results for Dy in OREAS 610 (Certified Value 1.53 ppm) | [24 empty] | Certified
Dy | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.02: 1.4; 2.03: 1.5; 2.04: 1.66; 2.05: 1.5; 2.06: 1.6 | 1
| 1 | 2 | 2.02: 1.4; 2.03: 1.5; 2.04: 1.65; 2.05: 1.5; 2.06: 1.55 | 10
| 1 | 3 | 2.02: 1.4; 2.03: 1.5; 2.04: 1.56; 2.05: 1.5; 2.06: 1.6 | 16
| 1 | 4 | 2.02: 1.4; 2.03: 1.5; 2.04: 1.64; 2.05: 1.6; 2.06: 1.55 | 1.525
| 1 | 5 | 2.02: 1.3; 2.03: 1.5; 2.04: 1.57; 2.05: 1.6; 2.06: 1.55 | 25
| 1 | 6 | 2.02: 1.4; 2.03: 1.5; 2.04: 1.62; 2.05: 1.6; 2.06: 1.6
| Mean |  | 2.02: 1.3833; 2.03: 1.5; 2.04: 1.6167; 2.05: 1.55; 2.06: 1.575
| Median |  | 2.02: 1.4; 2.03: 1.5; 2.04: 1.63; 2.05: 1.55; 2.06: 1.575
| Std Dev. |  | 2.02: 0.040825; 2.03: 0; 2.04: 0.042269; 2.05: 0.054772; 2.06: 0.027386
| Rel.Std.Dev. |  | 2.02: 2.951%; 2.03: 0%; 2.04: 2.615%; 2.05: 3.534%; 2.06: 1.739%
| PDM3 |  | 2.02: -9.29%; 2.03: -1.639%; 2.04: 6.011%; 2.05: 1.639%; 2.06: 3.279%
| Z-Score (Absolute) |  | 2.02: 2.25; 2.03: 0.67; 2.04: 0.9; 2.05: 0; 2.06: 0.34
| Analytical results for Er in OREAS 610 (Certified Value 0.59 ppm) | [24 empty] | Certified
Er | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.02: 0.5; 2.03: 0.6; 2.04: 0.65; 2.05: 0.5; 2.06: 0.6 | 1
| 1 | 2 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.65; 2.05: 0.6; 2.06: 0.6 | 11
| 1 | 3 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.62; 2.05: 0.5; 2.06: 0.6 | 16
| 1 | 4 | 2.02: 0.6; 2.03: 0.5; 2.04: 0.62; 2.05: 0.6; 2.06: 0.6 | 0.59367
| 1 | 5 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.64; 2.05: 0.6; 2.06: 0.65 | 26
| 1 | 6 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.63; 2.05: 0.6; 2.06: 0.6
| Mean |  | 2.02: 0.58333; 2.03: 0.58333; 2.04: 0.635; 2.05: 0.56667; 2.06: 0.60833
| Median |  | 2.02: 0.6; 2.03: 0.6; 2.04: 0.635; 2.05: 0.6; 2.06: 0.6
| Std Dev. |  | 2.02: 0.040825; 2.03: 0.040825; 2.04: 0.013784; 2.05: 0.05164; 2.06: 0.020412
| Rel.Std.Dev. |  | 2.02: 6.999%; 2.03: 6.999%; 2.04: 2.171%; 2.05: 9.113%; 2.06: 3.355%
| PDM3 |  | 2.02: -1.741%; 2.03: -1.741%; 2.04: 6.962%; 2.05: -4.548%; 2.06: 2.471%
| Z-Score (Absolute) |  | 2.02: 0; 2.03: 0; 2.04: 2.09; 2.05: 0.67; 2.06: 1.01
| Analytical results for Eu in OREAS 610 (Certified Value 0.77 ppm) | [24 empty] | Certified
Eu | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.02: 0.8; 2.03: 0.78; 2.04: 0.84; 2.05: 0.7; 2.06: 0.75 | 1
| 1 | 2 | 2.02: 0.8; 2.03: 0.74; 2.04: 0.83; 2.05: 0.7; 2.06: 0.75 | 12
| 1 | 3 | 2.02: 0.8; 2.03: 0.69; 2.04: 0.79; 2.05: 0.7; 2.06: 0.8 | 16
| 1 | 4 | 2.02: 0.7; 2.03: 0.73; 2.04: 0.82; 2.05: 0.7; 2.06: 0.75 | 0.77133
| 1 | 5 | 2.02: 0.8; 2.03: 0.76; 2.04: 0.81; 2.05: 0.8; 2.06: 0.8 | 27
| 1 | 6 | 2.02: 0.9; 2.03: 0.74; 2.04: 0.81; 2.05: 0.8; 2.06: 0.75
| Mean |  | 2.02: 0.8; 2.03: 0.74; 2.04: 0.81667; 2.05: 0.73333; 2.06: 0.76667
| Median |  | 2.02: 0.8; 2.03: 0.74; 2.04: 0.815; 2.05: 0.7; 2.06: 0.75
| Std Dev. |  | 2.02: 0.063246; 2.03: 0.030332; 2.04: 0.017512; 2.05: 0.05164; 2.06: 0.02582
| Rel.Std.Dev. |  | 2.02: 7.906%; 2.03: 4.099%; 2.04: 2.144%; 2.05: 7.042%; 2.06: 3.368%
| PDM3 |  | 2.02: 3.717%; 2.03: -4.062%; 2.04: 5.877%; 2.05: -4.927%; 2.06: -0.605%
| Z-Score (Absolute) |  | 2.02: 0.67; 2.03: 0.54; 2.04: 1.01; 2.05: 0.67; 2.06: 0
| Analytical results for Fe in OREAS 610 (Certified Value 2.37 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.38 | 2.25 | 2.31 | 2.26 | 2.38 | 2.2 | 2.28 | 2.37 | 2.39 | 2.55 | 2.4821 | 2.3261 | 2.42 | 2.51 | 2.3009 | 2.3 | 2.35 | 2.36 | 2.47 | 2.37 | 2.36 | 2.39 | 2.583 | 1
| 1 | 2 | 2.32 | 2.21 | 2.34 | 2.32 | 2.36 | 2.36 | 2.21 | 2.44 | 2.47 | 2.49 | 2.5128 | 2.3956 | 2.47 | 2.5 | 2.3247 | 2.25 | 2.272 | 2.41 | 2.53 | 2.47 | 2.35 | 2.38 | 2.558 | #N/A
| 1 | 3 | 2.38 | 2.25 | 2.3 | 2.3 | 2.39 | 2.33 | 2.22 | 2.33 | 2.42 | 2.51 | 2.5314 | 2.3928 | 2.32 | 2.58 | 2.3308 | 2.31 | 2.395 | 2.42 | 2.48 | 2.53 | 2.32 | 2.38 | 2.533 | 16
| 1 | 4 | 2.35 | 2.23 | 2.29 | 2.32 | 2.37 | 2.36 | 2.27 | 2.36 | 2.42 | 2.58 | 2.5164 | 2.3864 | 2.39 | 2.5 | 2.341 | 2.36 | 2.34 | 2.33 | 2.51 | 2.54 | 2.34 | 2.37 | 2.508 | 2.3658
| 1 | 5 | 2.39 | 2.24 | 2.42 | 2.28 | 2.36 | 2.36 | 2.2 | 2.34 | 2.42 | 2.56 | 2.4942 | 2.4913 | 2.49 | 2.54 | 2.3409 | 2.35 | 2.345 | 2.3 | 2.51 | 2.35 | 2.43 | 2.39 | 2.483 | 28
| 1 | 6 | 2.35 | 2.3 | 2.37 | 2.25 | 2.35 | 2.33 | 2.24 | 2.45 | 2.34 | 2.52 | 2.5192 | 2.4064 | 2.38 | 2.52 | 2.3215 | 2.4 | 2.387 | 2.33 | 2.42 | 2.4 | 2.39 | 2.36 | 2.558
| Mean |  | 2.3617 | 2.2467 | 2.3383 | 2.2883 | 2.3683 | 2.3233 | 2.2367 | 2.3817 | 2.41 | 2.535 | 2.5093 | 2.3998 | 2.4117 | 2.525 | 2.3266 | 2.3283 | 2.3482 | 2.3583 | 2.4867 | 2.4433 | 2.365 | 2.3783 | 2.5372
| Median |  | 2.365 | 2.245 | 2.325 | 2.29 | 2.365 | 2.345 | 2.23 | 2.365 | 2.42 | 2.535 | 2.5146 | 2.3942 | 2.405 | 2.515 | 2.3277 | 2.33 | 2.3475 | 2.345 | 2.495 | 2.435 | 2.355 | 2.38 | 2.5455
| Std Dev. |  | 0.026394 | 0.030111 | 0.049565 | 0.029944 | 0.01472 | 0.062183 | 0.03266 | 0.051153 | 0.042895 | 0.033912 | 0.017984 | 0.053092 | 0.062423 | 0.030822 | 0.014953 | 0.052694 | 0.043779 | 0.047924 | 0.039328 | 0.081894 | 0.03937 | 0.01169 | 0.036799
| Rel.Std.Dev. |  | 1.118% | 1.34% | 2.12% | 1.309% | 0.6215% | 2.676% | 1.46% | 2.148% | 1.78% | 1.338% | 0.7167% | 2.212% | 2.588% | 1.221% | 0.6427% | 2.263% | 1.864% | 2.032% | 1.582% | 3.352% | 1.665% | 0.4915% | 1.45%
| PDM3 |  | -0.1767% | -5.038% | -1.163% | -3.276% | 0.1051% | -1.797% | -5.46% | 0.6687% | 1.866% | 7.15% | 6.066% | 1.434% | 1.937% | 6.727% | -1.658% | -1.586% | -0.7473% | -0.3176% | 5.107% | 3.275% | -0.03579% | 0.5278% | 7.241%
| Z-Score (Absolute) |  | 0.11 | 1.97 | 0.49 | 1.29 | 0 | 0.73 | 2.13 | 0.22 | 0.67 | 2.69 | 2.28 | 0.51 | 0.7 | 2.53 | 0.67 | 0.65 | 0.33 | 0.16 | 1.91 | 1.21 | 0.05 | 0.16 | 2.73
| Analytical results for Ga in OREAS 610 (Certified Value 24.4 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 26.1 | 22.55 | 25.1 | 25 | 24.8 | 23.2 | 24.7 | 28.1 | 24.5 | 25.38 | 25.1422 | 19.48 | 27.5 | 23.4 | 25.0485 | 22.1 | 23.2 | 23.2 | 26 | 22 |  |  |  | 1
| 1 | 2 | 25.6 | 24.74 | 25.4 | 24.2 | 24.6 | 24.9 | 24.5 | 28.4 | 24.2 | 24.21 | 25.3658 | 23.11 | 26.9 | 24.2 | 24.7378 | 21 | 22.3 | 22.8 | 25 | 22 |  |  |  | 41
| 1 | 3 | 26.2 | 23.82 | 25.1 | 26 | 25.4 | 24.4 | 24 | 26.7 | 23.9 | 24.15 | 25.4669 | 22.63 | 27.5 | 24.6 | 24.2658 | 21.7 | 23.2 | 24 | 25 | 22 |  |  |  | 16
| 1 | 4 | 26.2 | 24.73 | 25.2 | 24 | 25 | 24.4 | 25.1 | 26.9 | 24.2 | 24.6 | 24.1451 | 21.33 | 27.3 | 24.4 | 23.542 | 22.2 | 22.5 | 22.9 | 26 | 22 |  |  |  | 24.419
| 1 | 5 | 26.5 | 24.4 | 26.4 | 24.7 | 24 | 24.7 | 23.6 | 25.5 | 25.1 | 24.93 | 24.0725 | 24 | 26.6 | 24.4 | 24.4563 | 22.3 | 22.4 | 22.6 | 25 | 22 |  |  |  | 29
| 1 | 6 | 26.2 | 24.32 | 25.3 | 24.4 | 25 | 24.3 | 23.8 | 27.5 | 25.1 | 25.3 | 24.7098 | 24 | 26.9 | 24.6 | 23.2088 | 21.5 | 22.3 | 23 | 25 | 22
| Mean |  | 26.1333 | 24.0933 | 25.4167 | 24.7167 | 24.8 | 24.3167 | 24.2833 | 27.1833 | 24.5 | 24.7617 | 24.817 | 22.425 | 27.1167 | 24.2667 | 24.2099 | 21.8 | 22.65 | 23.0833 | 25.3333 | 22
| Median |  | 26.2 | 24.36 | 25.25 | 24.55 | 24.9 | 24.4 | 24.25 | 27.2 | 24.35 | 24.765 | 24.926 | 22.87 | 27.1 | 24.4 | 24.361 | 21.9 | 22.45 | 22.95 | 25 | 22
| Std Dev. |  | 0.29439 | 0.82771 | 0.49565 | 0.72226 | 0.47329 | 0.59133 | 0.57764 | 1.0553 | 0.502 | 0.53011 | 0.6077 | 1.7518 | 0.37103 | 0.45019 | 0.70628 | 0.498 | 0.43243 | 0.4916 | 0.5164 | 0
| Rel.Std.Dev. |  | 1.127% | 3.435% | 1.95% | 2.922% | 1.908% | 2.432% | 2.379% | 3.882% | 2.049% | 2.141% | 2.449% | 7.812% | 1.368% | 1.855% | 2.917% | 2.284% | 1.909% | 2.13% | 2.038% | 0%
| PDM3 |  | 7.02% | -1.334% | 4.085% | 1.219% | 1.56% | -0.4192% | -0.5557% | 11.32% | 0.3316% | 1.403% | 1.63% | -8.166% | 11.05% | -0.624% | -0.8566% | -10.73% | -7.244% | -5.47% | 3.744% | -9.906%
| Z-Score (Absolute) |  | 1.74 | 0.32 | 1.02 | 0.31 | 0.4 | 0.09 | 0.13 | 2.81 | 0.09 | 0.36 | 0.41 | 2.01 | 2.74 | 0.14 | 0.2 | 2.64 | 1.78 | 1.34 | 0.94 | 2.44
| Analytical results for Gd in OREAS 610 (Certified Value 2.86 ppm) | [24 empty] | Certified
Gd | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 |  |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.02: 2.6; 2.03: 2.7; 2.04: 3.02; 2.05: 2.5; 2.06: 3.2 |  |  |  | 1
| 1 | 2 | 2.02: 2.6; 2.03: 2.7; 2.04: 3.1; 2.05: 2.6; 2.06: 3.2 |  |  |  | 13
| 1 | 3 | 2.02: 2.5; 2.03: 2.8; 2.04: 3.09; 2.05: 2.5; 2.06: 3.2 |  |  |  | 16
| 1 | 4 | 2.02: 2.8; 2.03: 2.8; 2.04: 3.05; 2.05: 2.7; 2.06: 3.2 |  |  |  | 2.856
| 1 | 5 | 2.02: 2.9; 2.03: 2.7; 2.04: 2.99; 2.05: 2.4; 2.06: 3.2 |  |  |  | 30
| 1 | 6 | 2.02: 2.8; 2.03: 2.7; 2.04: 3.13; 2.05: 2.8; 2.06: 3.2
| Mean |  | 2.02: 2.7; 2.03: 2.7333; 2.04: 3.0633; 2.05: 2.5833; 2.06: 3.2
| Median |  | 2.02: 2.7; 2.03: 2.7; 2.04: 3.07; 2.05: 2.55; 2.06: 3.2
| Std Dev. |  | 2.02: 0.15492; 2.03: 0.05164; 2.04: 0.052789; 2.05: 0.1472; 2.06: 4.8648e-16
| Rel.Std.Dev. |  | 2.02: 5.738%; 2.03: 1.889%; 2.04: 1.723%; 2.05: 5.698%; 2.06: 1.52e-14%
| PDM3 |  | 2.02: -5.462%; 2.03: -4.295%; 2.04: 7.26%; 2.05: -9.547%; 2.06: 12.04%
| Z-Score (Absolute) |  | 2.02: 0.15; 2.03: 0; 2.04: 1.48; 2.05: 0.67; 2.06: 2.1
| Analytical results for Ge in OREAS 610 (Indicative Value 2.13 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.18 | 2.20 | 2.23 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 5.9 | 2 | 0.15 | 0.18 | 0.1 | 0.14 | 0.16 | 5.3 | 5.6 | 3.8454 | 1.86 | 0.25 | 0.24 | 0.6 | 7.8 | [8 empty] | 1
| 1 | 2 | 5.4 | 1.9 | 0.18 | 0.18 | 0.2 | 0.12 | 0.16 | 5.4 | 5.6 | 3.7445 | 0.97 | 0.25 | 0.27 | 0.5 | 7.6 | [8 empty] | 6
| 1 | 3 | 5.3 | 1.7 | 0.15 | 0.21 | 0.19 | 0.13 | 0.16 | 5.4 | 5.7 | 3.8316 | 1.02 | 0.24 | 0.18 | 0.5 | 7.8 | [8 empty] | 16
| 1 | 4 | 5.2 | 2.3 | 0.16 | 0.21 | 0.18 | 0.12 | 0.14 | 5.3 | 5.5 | 3.7511 | 1.23 | 0.23 | 0.23 | 0.5 | 7.6 | [8 empty] | 2.1336
| 1 | 5 | 4.6 | 0.9 | 0.17 | 0.18 | 0.2 | 0.11 | 0.16 | 5.1 | 5.3 | 3.7749 | 0.24 | 0.26 | 0.15 | 0.4 | 7.6 | [8 empty] | 12
| 1 | 6 | 5.4 | 0.8 | 0.16 | 0.16 | 0.19 | 0.14 | 0.17 | 5.4 | 5.6 | 3.731 | 0.16 | 0.27 | 0.24 | 0.5 | 7.8
| Mean |  | 5.3 | 1.6 | 0.16167 | 0.18667 | 0.17667 | 0.12667 | 0.15833 | 5.3167 | 5.55 | 3.7797 | 0.91333 | 0.25 | 0.21833 | 0.5 | 7.7
| Median |  | 5.35 | 1.8 | 0.16 | 0.18 | 0.19 | 0.125 | 0.16 | 5.35 | 5.6 | 3.763 | 0.995 | 0.25 | 0.235 | 0.5 | 7.7
| Std Dev. |  | 0.41952 | 0.61319 | 0.01169 | 0.019664 | 0.038297 | 0.012111 | 0.0098319 | 0.1169 | 0.13784 | 0.047904 | 0.63748 | 0.014142 | 0.04446 | 0.063246 | 0.10954
| Rel.Std.Dev. |  | 7.916% | 38.32% | 7.231% | 10.53% | 21.68% | 9.561% | 6.21% | 2.199% | 2.484% | 1.267% | 69.8% | 5.657% | 20.36% | 12.65% | 1.423%
| PDM3 |  | 148.4% | -25.01% | -92.42% | -91.25% | -91.72% | -94.06% | -92.58% | 149.2% | 160.1% | 77.16% | -57.19% | -88.28% | -89.77% | -76.56% | 260.9%
| Z-Score (Absolute) |  | 8.67 | 1.99 | 0.61 | 0.57 | 0.58 | 0.67 | 0.62 | 8.7 | 9.12 | 5.92 | 0.75 | 0.45 | 0.51 | 0 | 13
| Analytical results for Hf in OREAS 610 (Certified Value 2.01 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.8 | 1.72 | 2.2 | 1.8 | 2 | 1.9 | 2 | 1.9 | 2 | 2.17 | 2.0202 | 2.03 | 2 | 2.4 | 1.78 | 2.21 | 1.94 | 2.2 | [5 empty] | 1
| 1 | 2 | 1.8 | 1.87 | 2.2 | 1.9 | 2.1 | 2 | 2 | 2 | 2 | 2.13 | 2.0411 | 2.1 | 2 | 2.4 | 1.91 | 2.2 | 1.91 | 2.2 | [5 empty] | 27
| 1 | 3 | 1.8 | 1.83 | 2.1 | 2.1 | 2.1 | 1.9 | 2 | 1.9 | 2 | 2.12 | 2.0731 | 2.08 | 2 | 2.4 | 1.76 | 2.2 | 2.02 | 2 | [5 empty] | 16
| 1 | 4 | 1.8 | 1.81 | 2.2 | 1.9 | 2.1 | 2 | 1.9 | 1.9 | 2.1 | 2.15 | 2.0198 | 2.17 | 2 | 2.4 | 1.87 | 2.15 | 1.94 | 2.1 | [5 empty] | 2.0068
| 1 | 5 | 1.8 | 1.85 | 2.2 | 2 | 2 | 2 | 2 | 1.9 | 2.1 | 2.15 | 2.0346 | 2.1 | 2 | 2.6 | 1.91 | 2.13 | 1.87 | 2.1 | [5 empty] | 31
| 1 | 6 | 1.8 | 1.92 | 2.2 | 1.9 | 2 | 2 | 2 | 2 | 2 | 2.13 | 2.0656 | 2.25 | 2 | 2.6 | 1.83 | 2.21 | 1.92 | 2.2
| Mean |  | 1.8 | 1.8333 | 2.1833 | 1.9333 | 2.05 | 1.9667 | 1.9833 | 1.9333 | 2.0333 | 2.1417 | 2.0424 | 2.1217 | 2 | 2.4667 | 1.8433 | 2.1833 | 1.9333 | 2.1333
| Median |  | 1.8 | 1.84 | 2.2 | 1.9 | 2.05 | 2 | 2 | 1.9 | 2 | 2.14 | 2.0378 | 2.1 | 2 | 2.4 | 1.85 | 2.2 | 1.93 | 2.15
| Std Dev. |  | 0 | 0.067132 | 0.040825 | 0.10328 | 0.054772 | 0.05164 | 0.040825 | 0.05164 | 0.05164 | 0.018348 | 0.02257 | 0.077309 | 0 | 0.10328 | 0.064395 | 0.034448 | 0.049666 | 0.08165
| Rel.Std.Dev. |  | 0% | 3.662% | 1.87% | 5.342% | 2.672% | 2.626% | 2.058% | 2.671% | 2.54% | 0.8567% | 1.105% | 3.644% | 0% | 4.187% | 3.493% | 1.578% | 2.569% | 3.827%
| PDM3 |  | -10.31% | -8.644% | 8.796% | -3.661% | 2.152% | -2% | -1.17% | -3.661% | 1.322% | 6.72% | 1.773% | 5.724% | -0.3392% | 22.92% | -8.146% | 8.796% | -3.661% | 6.305%
| Z-Score (Absolute) |  | 1.55 | 1.31 | 1.19 | 0.6 | 0.24 | 0.36 | 0.24 | 0.6 | 0.12 | 0.9 | 0.18 | 0.75 | 0.12 | 3.22 | 1.24 | 1.19 | 0.6 | 0.84
Results from laboratory 15 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 610 (Indicative Value 0.43 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | 2.02: Lab; 2.03: Lab | [5 empty] | 1
| No | No | 2.02: 2.04; 2.03: 2.17 | [5 empty] | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 2.02: 0.6; 2.03: < 2 | [5 empty] | 1
| 1 | 2 | 2.02: 0.11; 2.03: < 2 | [5 empty] | 5
| 1 | 3 | 2.02: 0.02; 2.03: < 2 | [5 empty] | 16
| 1 | 4 | 2.02: 0.27; 2.03: < 2 | [5 empty] | 0.42833
| 1 | 5 | 2.02: 0.78; 2.03: < 2 | [5 empty] | 11
| 1 | 6 | 2.02: 0.79; 2.03: < 2
| Mean |  | 2.02: 0.42833
| Median |  | 2.02: 0.435
| Std Dev. |  | 2.02: 0.33973
| Rel.Std.Dev. |  | 2.02: 79.31%
| PDM3 |  | 2.02: 6.661e-14%
| Z-Score (Absolute) |  | 2.02: 0.67; 2.03: 0.67
| Analytical results for Ho in OREAS 610 (Certified Value 0.22 ppm) | [24 empty] | Certified
Ho | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | [5 empty] | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | [5 empty] | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.24; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 8
| 1 | 3 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 0.217
| 1 | 5 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.24; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 32
| 1 | 6 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.24; 2.05: 0.2; 2.06: 0.24
| Mean |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.245; 2.05: 0.2; 2.06: 0.24
| Median |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.245; 2.05: 0.2; 2.06: 0.24
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 3.0405e-17; 2.04: 0.0054772; 2.05: 3.0405e-17; 2.06: 0
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 1.52e-14%; 2.04: 2.236%; 2.05: 1.52e-14%; 2.06: 0%
| PDM3 |  | 2.02: -7.834%; 2.03: -7.834%; 2.04: 12.9%; 2.05: -7.834%; 2.06: 10.6%
Z-Score (Absolute) | bar: 2.02 to 2.06
| Analytical results for In in OREAS 610 (Certified Value 3.9 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 3.44 | 4.04 | 3.7 | 3.6 | 3.82 | 3.69 | 3.93 | 4.19 | 4.07 | 4.01 | 4.0291 | 4.12 | 3.8 | 3.88 | 4.8 | 3.8 | 3.55 | 4.11 | [5 empty] | 1
| 1 | 2 | 3.43 | 4.23 | 3.6 | 3.58 | 3.85 | 3.95 | 3.88 | 4.33 | 4.08 | 3.95 | 4.0379 | 4.35 | 3.68 | 3.84 | 4.6 | 3.89 | 3.56 | 4.06 | [5 empty] | 29
| 1 | 3 | 3.52 | 4.18 | 3.7 | 3.75 | 3.93 | 3.8 | 3.8 | 4.1 | 4.03 | 4.01 | 4.0924 | 4.17 | 3.72 | 3.88 | 4.6 | 3.9 | 3.64 | 4.01 | [5 empty] | 16
| 1 | 4 | 3.46 | 4.52 | 3.5 | 3.62 | 3.89 | 3.87 | 3.96 | 4.23 | 4.15 | 3.92 | 4.0477 | 4.32 | 3.87 | 3.94 | 4.8 | 3.72 | 3.51 | 3.99 | [5 empty] | 3.8952
| 1 | 5 | 3.53 | 4.2 | 3.8 | 3.74 | 3.63 | 3.86 | 3.7 | 4.06 | 4.22 | 3.98 | 4.1112 | 4.22 | 3.68 | 3.86 | 4.7 | 3.8 | 3.51 | 3.96 | [5 empty] | 33
| 1 | 6 | 3.5 | 4.44 | 3.6 | 3.71 | 3.85 | 3.84 | 3.77 | 4.22 | 4.24 | 4.02 | 4.1228 | 4.28 | 3.72 | 3.86 | 4.4 | 3.71 | 3.53 | 3.97
| Mean |  | 3.48 | 4.2683 | 3.65 | 3.6667 | 3.8283 | 3.835 | 3.84 | 4.1883 | 4.1317 | 3.9817 | 4.0735 | 4.2433 | 3.745 | 3.8767 | 4.65 | 3.8033 | 3.55 | 4.0167
| Median |  | 3.48 | 4.215 | 3.65 | 3.665 | 3.85 | 3.85 | 3.84 | 4.205 | 4.115 | 3.995 | 4.07 | 4.25 | 3.72 | 3.87 | 4.65 | 3.8 | 3.54 | 4
| Std Dev. |  | 0.042426 | 0.17826 | 0.10488 | 0.075277 | 0.10439 | 0.086429 | 0.1002 | 0.09704 | 0.085654 | 0.039707 | 0.040293 | 0.089144 | 0.075299 | 0.034448 | 0.15166 | 0.080664 | 0.04858 | 0.05785
| Rel.Std.Dev. |  | 1.219% | 4.176% | 2.873% | 2.053% | 2.727% | 2.254% | 2.609% | 2.317% | 2.073% | 0.9973% | 0.9891% | 2.101% | 2.011% | 0.8886% | 3.261% | 2.121% | 1.368% | 1.44%
| PDM3 |  | -10.66% | 9.58% | -6.295% | -5.867% | -1.716% | -1.545% | -1.417% | 7.526% | 6.071% | 2.22% | 4.579% | 8.938% | -3.856% | -0.4755% | 19.38% | -2.358% | -8.862% | 3.119%
| Z-Score (Absolute) |  | 1.28 | 1.38 | 0.7 | 0.65 | 0.1 | 0.08 | 0.06 | 1.11 | 0.92 | 0.42 | 0.73 | 1.3 | 0.38 | 0.06 | 2.67 | 0.19 | 1.04 | 0.53
| Analytical results for K in OREAS 610 (Certified Value 1.98 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 1.953 | 2.36 | 1.94 | 1.9 | 1.96 | 1.85 | 1.91 | 1.97 | 2.03 | 2.061 | 2.0073 | 1.979 | 2.02 | 1.91 | 1.9868 | 1.9 | 1.924 | 1.91 | 2 | 1.89 | 2.09 | 2.1 | 2.027 | 1
| 1 | 2 | 1.9173 | 2.43 | 2 | 1.94 | 1.94 | 1.94 | 1.88 | 2.04 | 2.09 | 2.0163 | 2.0236 | 2.0411 | 2.11 | 1.94 | 2.0189 | 1.86 | 1.864 | 1.93 | 2.06 | 1.93 | 2.09 | 2.09 | 2.013 | #N/A
| 1 | 3 | 1.9513 | 2.35 | 1.94 | 1.93 | 1.97 | 1.92 | 1.87 | 1.93 | 2.07 | 2.0133 | 2.0023 | 2.0172 | 1.92 | 2 | 2.0338 | 1.88 | 1.896 | 1.96 | 1.99 | 1.96 | 2.08 | 2.1 | 1.999 | 16
| 1 | 4 | 1.9309 | 2.33 | 1.95 | 1.95 | 1.95 | 1.93 | 1.9 | 1.98 | 2.06 | 2.0667 | 2.0212 | 2.0083 | 2.02 | 1.93 | 2.0212 | 1.95 | 1.9 | 1.91 | 2.03 | 1.97 | 2.1 | 2.1 | 1.985 | 1.9786
| 1 | 5 | 1.9436 | 2.4 | 1.91 | 1.93 | 1.95 | 1.94 | 1.86 | 1.93 | 2.06 | 2.0636 | 2.0089 | 2.0399 | 2.09 | 2 | 2.0025 | 1.93 | 1.891 | 1.88 | 2.04 | 1.88 | 2.11 | 2.09 | 1.991 | 34
| 1 | 6 | 1.9334 | 2.44 | 1.85 | 1.92 | 1.93 | 1.91 | 1.88 | 2.05 | 1.98 | 2.0309 | 2.0277 | 2.0253 | 1.95 | 1.97 | 1.9973 | 1.96 | 1.918 | 1.9 | 2.01 | 1.92 | 2.08 | 2.1 | 1.987
| Mean |  | 1.9383 | 2.385 | 1.9317 | 1.9283 | 1.95 | 1.915 | 1.8833 | 1.9833 | 2.0483 | 2.042 | 2.0152 | 2.0185 | 2.0183 | 1.9583 | 2.0101 | 1.9133 | 1.8988 | 1.915 | 2.0217 | 1.925 | 2.0917 | 2.0967 | 2.0003
| Median |  | 1.9385 | 2.38 | 1.94 | 1.93 | 1.95 | 1.925 | 1.88 | 1.975 | 2.06 | 2.0459 | 2.015 | 2.0213 | 2.02 | 1.955 | 2.0107 | 1.915 | 1.898 | 1.91 | 2.02 | 1.925 | 2.09 | 2.1 | 1.995
| Std Dev. |  | 0.01366 | 0.045056 | 0.049565 | 0.017224 | 0.014142 | 0.033912 | 0.018619 | 0.052026 | 0.038687 | 0.024678 | 0.010307 | 0.023156 | 0.074677 | 0.037639 | 0.017449 | 0.039833 | 0.021377 | 0.027386 | 0.026394 | 0.036194 | 0.01169 | 0.005164 | 0.016573
| Rel.Std.Dev. |  | 0.7047% | 1.889% | 2.566% | 0.8932% | 0.7252% | 1.771% | 0.9886% | 2.623% | 1.889% | 1.209% | 0.5115% | 1.147% | 3.7% | 1.922% | 0.8681% | 2.082% | 1.126% | 1.43% | 1.306% | 1.88% | 0.5589% | 0.2463% | 0.8285%
| PDM3 |  | -2.041% | 20.54% | -2.373% | -2.542% | -1.447% | -3.216% | -4.816% | 0.2379% | 3.523% | 3.201% | 1.847% | 2.014% | 2.007% | -1.026% | 1.59% | -3.3% | -4.033% | -3.216% | 2.175% | -2.71% | 5.713% | 5.966% | 1.097%
| Z-Score (Absolute) |  | 0.55 | 4.92 | 0.63 | 0.67 | 0.41 | 0.84 | 1.23 | 0 | 0.8 | 0.72 | 0.39 | 0.43 | 0.43 | 0.31 | 0.33 | 0.86 | 1.04 | 0.84 | 0.47 | 0.72 | 1.33 | 1.39 | 0.21
| Analytical results for La in OREAS 610 (Certified Value 20.4 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 1
1
| 1 | 1 | 14.8 | 11 | 19.4 | 19.6 | 23.9 | 22.4 | 22.6 | 24.27 | 14.3 | 23.8 | 28.4 | 26.0307 | 16.3 | 19.386 | 16.3 | 21.1 | 17.3 | 25 | 17 | 25.24 |  |  |  | 1
| 1 | 2 | 15.1 | 8.7 | 19.9 | 19.7 | 21.9 | 21.4 | 23.3 | 23.62 | 15.2 | 23.4 | 26.8 | 26.7108 | 19.5 | 19.816 | 17.8 | 19.6 | 16.3 | 25 | 17 | 25.13 |  |  |  | 30
| 1 | 3 | 15.7 | 8.5 | 21.2 | 20.9 | 22 | 21.5 | 23.4 | 24.58 | 13.5 | 23.9 | 27.8 | 26.4332 | 19.1 | 19.191 | 17.1 | 19.6 | 16.3 | 25 | 17 | 25.02 |  |  |  | 16
| 1 | 4 | 15.2 | 9.5 | 20.1 | 20.6 | 22.6 | 21.7 | 22.6 | 24.29 | 15.1 | 21.7 | 27.9 | 27.2397 | 18.1 | 17.498 | 16.8 | 20.6 | 17.7 | 24 | 17 | 24.91 |  |  |  | 20.3844
| 1 | 5 | 14.4 | 11.7 | 20.6 | 19.7 | 22.9 | 20.3 | 23.9 | 24.4 | 16.3 | 22.2 | 27.6 | 26.6557 | 19.2 | 18.145 | 16.6 | 19.2 | 16.4 | 24 | 17 | 24.8 |  |  |  | 35
| 1 | 6 | 14.2 | 11.9 | 21.8 | 20.1 | 21.9 | 20.8 | 23.9 | 24.96 | 15 | 22.1 | 28.4 | 27.0808 | 19 | 18.849 | 16.2 | 20.9 | 16.8 | 25 | 17 | 24.69
| Mean |  | 14.9 | 10.2167 | 20.5 | 20.1 | 22.5333 | 21.35 | 23.2833 | 24.3533 | 14.9 | 22.85 | 27.8167 | 26.6918 | 18.5333 | 18.8142 | 16.8 | 20.1667 | 16.8 | 24.6667 | 17 | 24.965
| Median |  | 14.95 | 10.25 | 20.35 | 19.9 | 22.3 | 21.45 | 23.35 | 24.345 | 15.05 | 22.8 | 27.85 | 26.6832 | 19.05 | 19.02 | 16.7 | 20.1 | 16.6 | 25 | 17 | 24.965
| Std Dev. |  | 0.55136 | 1.5105 | 0.88544 | 0.54037 | 0.78655 | 0.7287 | 0.58452 | 0.44053 | 0.94021 | 0.96073 | 0.5947 | 0.43752 | 1.1911 | 0.85486 | 0.58992 | 0.79666 | 0.58652 | 0.5164 | 0 | 0.20579
| Rel.Std.Dev. |  | 3.7% | 14.78% | 4.319% | 2.688% | 3.491% | 3.413% | 2.51% | 1.809% | 6.31% | 4.205% | 2.138% | 1.639% | 6.427% | 4.544% | 3.511% | 3.95% | 3.491% | 2.094% | 0% | 0.8243%
| PDM3 |  | -26.9% | -49.88% | 0.5672% | -1.395% | 10.54% | 4.737% | 14.22% | 19.47% | -26.9% | 12.1% | 36.46% | 30.94% | -9.081% | -7.703% | -17.58% | -1.068% | -17.58% | 21.01% | -16.6% | 22.47%
| Z-Score (Absolute) |  | 1.07 | 1.99 | 0.03 | 0.05 | 0.43 | 0.2 | 0.58 | 0.79 | 1.07 | 0.49 | 1.47 | 1.25 | 0.35 | 0.3 | 0.69 | 0.03 | 0.69 | 0.85 | 0.65 | 0.91
| Analytical results for Li in OREAS 610 (Certified Value 28.4 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 26.9 | 26.3 | 28 | 27.9 | 28.3 | 26.4 | 30.1 | 32.6 | 29.4 | 28.8 | 28.5647 | 30.4 | 28.8 | 28 | 31.6987 | 28.1 | 25.359 | 28 | 28 | 28 | 37.1 | 27 | 28.88 | 1
| 1 | 2 | 27.1 | 27 | 27.9 | 27.9 | 28.2 | 30.3 | 30.4 | 33.5 | 28.1 | 28.2 | 28.873 | 31 | 29.9 | 28.5 | 32.6397 | 26.9 | 25.333 | 31 | 29 | 29 | 39.2 | 27 | 28.54 | #N/A
| 1 | 3 | 27.8 | 27.6 | 26.7 | 28.9 | 28.8 | 29.1 | 29.4 | 31.1 | 29 | 27.5 | 28.9105 | 31.3 | 27.1 | 29.5 | 31.7907 | 27.6 | 26.442 | 30 | 28 | 29 | 37.3 | 27 | 28.2 | 16
| 1 | 4 | 28.8 | 30.3 | 27.1 | 26.3 | 28.7 | 29.6 | 30.7 | 32 | 28.7 | 27.2 | 27.6268 | 30.7 | 28.7 | 29.5 | 31.7081 | 27.1 | 25.486 | 30 | 28 | 29 | 30.4 | 27 | 27.86 | 28.4423
| 1 | 5 | 29 | 30.7 | 28.8 | 29.3 | 27.7 | 29.8 | 29.1 | 30.2 | 28.2 | 27.2 | 28.5212 | 31 | 29.7 | 29.5 | 31.3318 | 27.8 | 25.991 | 30 | 29 | 28 | 35.4 | 27 | 27.52 | 36
| 1 | 6 | 28.2 | 31.5 | 28 | 26.6 | 28.3 | 29.3 | 29.4 | 31.5 | 28.1 | 26.9 | 27.7015 | 31 | 27.7 | 29.5 | 31.6082 | 28.9 | 25.065 | 29 | 28 | 29 | 32.9 | 27 | 27.98
| Mean |  | 27.9667 | 28.9 | 27.75 | 27.8167 | 28.3333 | 29.0833 | 29.85 | 31.8167 | 28.5833 | 27.6333 | 28.3663 | 30.9 | 28.65 | 29.0833 | 31.7962 | 27.7333 | 25.6127 | 29.6667 | 28.3333 | 28.6667 | 35.3833 | 27 | 28.1633
| Median |  | 28 | 28.95 | 27.95 | 27.9 | 28.3 | 29.45 | 29.75 | 31.75 | 28.45 | 27.35 | 28.543 | 31 | 28.75 | 29.5 | 31.7034 | 27.7 | 25.4225 | 30 | 28 | 29 | 36.25 | 27 | 28.09
| Std Dev. |  | 0.8641 | 2.1918 | 0.74498 | 1.1974 | 0.39328 | 1.379 | 0.64109 | 1.1583 | 0.54191 | 0.72296 | 0.56661 | 0.30984 | 1.095 | 0.66458 | 0.44267 | 0.72296 | 0.50783 | 1.0328 | 0.5164 | 0.5164 | 3.2307 | 0 | 0.48915
| Rel.Std.Dev. |  | 3.09% | 7.584% | 2.685% | 4.304% | 1.388% | 4.742% | 2.148% | 3.641% | 1.896% | 2.616% | 1.997% | 1.003% | 3.822% | 2.285% | 1.392% | 2.607% | 1.983% | 3.481% | 1.823% | 1.801% | 9.131% | 0% | 1.737%
| PDM3 |  | -1.672% | 1.609% | -2.434% | -2.2% | -0.383% | 2.254% | 4.949% | 11.86% | 0.4959% | -2.844% | -0.2672% | 8.641% | 0.7303% | 2.254% | 11.79% | -2.493% | -9.949% | 4.305% | -0.383% | 0.7889% | 24.4% | -5.071% | -0.9808%
| Z-Score (Absolute) |  | 0.54 | 0.28 | 0.73 | 0.67 | 0.22 | 0.44 | 1.11 | 2.84 | 0 | 0.84 | 0.19 | 2.04 | 0.06 | 0.44 | 2.83 | 0.75 | 2.61 | 0.95 | 0.22 | 0.07 | 5.98 | 1.39 | 0.37
| Analytical results for Lu in OREAS 610 (Indicative Value 80 ppb) | [24 empty] | Indicative
Lu | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16; 2.07: 2.20; 2.08: 2.21 | ppb
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*OES/MS | 1
1
| 1 | 1 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100; 2.07: 70; 2.08: 90 | 1
| 1 | 2 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100; 2.07: 80; 2.08: 90 | 6
| 1 | 3 | 2.02: 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100; 2.07: 80; 2.08: 80 | 16
| 1 | 4 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100; 2.07: 70; 2.08: 80 | 80
| 1 | 5 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100; 2.07: 70; 2.08: 80 | 12
| 1 | 6 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100; 2.07: 70; 2.08: 80
| Mean |  | 2.02: 100; 2.04: 85; 2.06: 100; 2.07: 73.3333; 2.08: 83.3333
| Median |  | 2.02: 100; 2.04: 85; 2.06: 100; 2.07: 70; 2.08: 80
| Std Dev. |  | 2.04: 5.4772; 2.06: 0; 2.07: 5.164; 2.08: 5.164
| Rel.Std.Dev. |  | 2.04: 6.444%; 2.06: 0%; 2.07: 7.042%; 2.08: 6.197%
| PDM3 |  | 2.02: 25%; 2.04: 6.25%; 2.06: 25%; 2.07: -8.333%; 2.08: 4.167%
| Z-Score (Absolute) |  | 2.02: 0.67; 2.03: 1.05; 2.04: 0.52; 2.05: 1.05; 2.06: 1.2; 2.07: 0; 2.08: 0.45
| Analytical results for Mg in OREAS 610 (Certified Value 1557 ppm) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 1575 | 1500 | 1600 | 1400 | 1500 | 1400 | 1500 | 1500 | 1690 | 1651 | 1601.4274 | 1461 | 1700 | 1600 | 1678.482 | 1600 | 1380 | 1500 | 1600 | 1500 | 1600 | 1100 | 1660 | 1
| 1 | 2 | 1537 | 1500 | 1500 | 1500 | 1500 | 1500 | 1400 | 1500 | 1710 | 1600 | 1604.6131 | 1578 | 1700 | 1500 | 1704.3731 | 1500 | 1360 | 1600 | 1600 | 1500 | 1600 | 1100 | 1644 | #N/A
| 1 | 3 | 1574 | 1500 | 1500 | 1400 | 1500 | 1500 | 1400 | 1500 | 1660 | 1605 | 1616.0701 | 1569 | 1700 | 1600 | 1711.2312 | 1600 | 1410 | 1500 | 1600 | 1600 | 1600 | 1000 | 1628 | 16
| 1 | 4 | 1559 | 1500 | 1500 | 1500 | 1500 | 1500 | 1500 | 1500 | 1670 | 1657 | 1624.9759 | 1578 | 1700 | 1600 | 1713.7657 | 1600 | 1330 | 1400 | 1600 | 1600 | 1600 | 1100 | 1612 | 1557.4306
| 1 | 5 | 1603 | 1600 | 1700 | 1400 | 1500 | 1500 | 1400 | 1500 | 1650 | 1636 | 1591.1499 | 1637 | 1700 | 1600 | 1697.5135 | 1600 | 1350 | 1500 | 1600 | 1500 | 1600 | 1000 | 1596 | 37
| 1 | 6 | 1564 | 1500 | 1600 | 1400 | 1500 | 1500 | 1500 | 1600 | 1620 | 1592 | 1601.2569 | 1536 | 1700 | 1600 | 1683.982 | 1600 | 1390 | 1400 | 1600 | 1500 | 1600 | 1100 | 1580
| Mean |  | 1568.6667 | 1516.6667 | 1566.6667 | 1433.3333 | 1500 | 1483.3333 | 1450 | 1516.6667 | 1666.6667 | 1623.5 | 1606.5822 | 1559.8333 | 1700 | 1583.3333 | 1698.2246 | 1583.3333 | 1370 | 1483.3333 | 1600 | 1533.3333 | 1600 | 1066.6667 | 1620
| Median |  | 1569 | 1500 | 1550 | 1400 | 1500 | 1500 | 1450 | 1500 | 1665 | 1620.5 | 1603.0203 | 1573.5 | 1700 | 1600 | 1700.9433 | 1600 | 1370 | 1500 | 1600 | 1500 | 1600 | 1100 | 1620
| Std Dev. |  | 21.7501 | 40.8248 | 81.6497 | 51.6398 | 0 | 40.8248 | 54.7723 | 40.8248 | 31.4113 | 28.0054 | 12.0442 | 58.3692 | 0 | 40.8248 | 14.4335 | 40.8248 | 28.9828 | 75.2773 | 0 | 51.6398 | 0 | 51.6398 | 29.9333
| Rel.Std.Dev. |  | 1.387% | 2.692% | 5.212% | 3.603% | 0% | 2.752% | 3.777% | 2.692% | 1.885% | 1.725% | 0.7497% | 3.742% | 0% | 2.578% | 0.8499% | 2.578% | 2.116% | 5.075% | 0% | 3.368% | 0% | 4.841% | 1.848%
| PDM3 |  | 0.7214% | -2.617% | 0.593% | -7.968% | -3.688% | -4.758% | -6.898% | -2.617% | 7.014% | 4.242% | 3.156% | 0.1543% | 9.154% | 1.663% | 9.04% | 1.663% | -12.03% | -4.758% | 2.733% | -1.547% | 2.733% | -31.51% | 4.017%
| Z-Score (Absolute) |  | 0.03 | 0.63 | 0 | 1.69 | 0.84 | 1.05 | 1.48 | 0.63 | 1.26 | 0.72 | 0.5 | 0.09 | 1.69 | 0.21 | 1.66 | 0.21 | 2.49 | 1.05 | 0.42 | 0.42 | 0.42 | 6.32 | 0.67
| Analytical results for Mn in OREAS 610 (Certified Value 78 ppm) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 76 | 84 | 76 | 75 | 76 | 74 | 77 | 78 | 76 | 84 | 78.2881 | 73 | 80 | 70 | 77.6469 | 74 | 71.973 | 71 | 82 | 75 | 83 | 93 | 75.25 | 1
| 1 | 2 | 75 | 86 | 71 | 78 | 78 | 77 | 80 | 80 | 77 | 81 | 76.7102 | 76 | 81 | 88 | 78.0547 | 74 | 71.139 | 71 | 84 | 75 | 81 | 93 | 77.19 | 32
| 1 | 3 | 77 | 82 | 66 | 78 | 79 | 76 | 76 | 78 | 74 | 82 | 78.8165 | 76 | 79 | 84 | 78.6905 | 76 | 71.481 | 73 | 83 | 76 | 82 | 93 | 76.13 | 16
| 1 | 4 | 75 | 83 | 78 | 78 | 77 | 77 | 77 | 78 | 77 | 85 | 76.7603 | 75 | 79 | 80 | 78.1732 | 76 | 73.255 | 71 | 84 | 75 | 80 | 93 | 79.07 | 77.7201
| 1 | 5 | 81 | 79 | 82 | 76 | 78 | 77 | 76 | 77 | 77 | 82 | 75.7396 | 81 | 81 | 78 | 78.1662 | 77 | 72.633 | 72 | 83 | 72 | 82 | 93 | 78.01 | 38
| 1 | 6 | 75 | 83 | 82 | 75 | 78 | 76 | 76 | 80 | 76 | 82 | 79.0795 | 78 | 79 | 82 | 77.2805 | 78 | 74.566 | 71 | 82 | 75 | 81 | 93 | 79.95
| Mean |  | 76.5 | 82.8333 | 75.8333 | 76.6667 | 77.6667 | 76.1667 | 77 | 78.5 | 76.1667 | 82.6667 | 77.5657 | 76.5 | 79.8333 | 80.3333 | 78.002 | 75.8333 | 72.5078 | 71.5 | 83 | 74.6667 | 81.5 | 93 | 77.6
| Median |  | 75.5 | 83 | 77 | 77 | 78 | 76.5 | 76.5 | 78 | 76.5 | 82 | 77.5242 | 76 | 79.5 | 81 | 78.1104 | 76 | 72.303 | 71 | 83 | 75 | 81.5 | 93 | 77.6
| Std Dev. |  | 2.3452 | 2.3166 | 6.3377 | 1.5055 | 1.0328 | 1.169 | 1.5492 | 1.2247 | 1.169 | 1.5055 | 1.3486 | 2.7386 | 0.98319 | 6.121 | 0.48575 | 1.6021 | 1.2668 | 0.83666 | 0.89443 | 1.3663 | 1.0488 | 0 | 1.7722
| Rel.Std.Dev. |  | 3.066% | 2.797% | 8.357% | 1.964% | 1.33% | 1.535% | 2.012% | 1.56% | 1.535% | 1.821% | 1.739% | 3.58% | 1.232% | 7.62% | 0.6227% | 2.113% | 1.747% | 1.17% | 1.078% | 1.83% | 1.287% | 0% | 2.284%
| PDM3 |  | -1.57% | 6.579% | -2.428% | -1.355% | -0.06875% | -1.999% | -0.9265% | 1.003% | -1.999% | 6.365% | -0.1987% | -1.57% | 2.719% | 3.362% | 0.3627% | -2.428% | -6.706% | -8.003% | 6.793% | -3.929% | 4.863% | 19.66% | -0.1545%
| Z-Score (Absolute) |  | 0.41 | 2.05 | 0.67 | 0.35 | 0.04 | 0.54 | 0.22 | 0.36 | 0.54 | 1.99 | 0 | 0.41 | 0.88 | 1.08 | 0.17 | 0.67 | 1.97 | 2.36 | 2.12 | 1.13 | 1.53 | 6.01 | 0.01
| Analytical results for Mo in OREAS 610 (Certified Value 4.82 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 4.4 | 4.62 | 4.62 | 4.83 | 4.81 | 4.61 | 4.7 | 4.95 | 4.7 | 5.1 | 5.0208 | 4.66 | 5.75 | 5 | 7.5206 | 4.06 | 4.406 | 5.04 | 4.51 | 4 | 5.1 | 5 | 8.977 | 1
| 1 | 2 | 4.4 | 4.67 | 4.54 | 4.87 | 4.77 | 4.99 | 5.07 | 5.64 | 4.9 | 5 | 4.9026 | 4.36 | 5.55 | 5 | 7.6359 | 3.91 | 5.099 | 5.27 | 4.62 | 4 | 5.5 | 5 | 8.24 | 33
| 1 | 3 | 4.6 | 4.62 | 4.61 | 4.9 | 5.18 | 5.08 | 5.24 | 4.84 | 4.7 | 4.8 | 4.9512 | 4.33 | 5.88 | 5 | 6.5576 | 4.19 | 4.829 | 5.22 | 4.71 | 4 | 5.2 | 4 | 8.503 | 16
| 1 | 4 | 4.6 | 4.86 | 4.41 | 4.67 | 4.84 | 5 | 5.62 | 5.37 | 4.7 | 4.9 | 5.0309 | 4.17 | 5.52 | 5 | 8.0133 | 4.34 | 4.807 | 5.03 | 4.54 | 4 | 5.3 | 4 | 7.966 | 4.8165
| 1 | 5 | 4.7 | 4.72 | 4.85 | 5.26 | 4.84 | 5.03 | 4.63 | 4.75 | 4.7 | 4.8 | 4.9127 | 5.15 | 5.8 | 6 | 7.2041 | 4.27 | 4.439 | 5.18 | 4.28 | 4 | 5.3 | 4 | 8.029 | 39
| 1 | 6 | 4.5 | 4.84 | 4.53 | 4.9 | 4.94 | 5.15 | 4.81 | 5.06 | 4.8 | 4.9 | 4.9472 | 5.46 | 5.61 | 5 | 7.0224 | 4.3 | 5.024 | 5.12 | 4.48 | 4 | 5.2 | 4 | 7.792
| Mean |  | 4.5333 | 4.7217 | 4.5933 | 4.905 | 4.8967 | 4.9767 | 5.0117 | 5.1017 | 4.75 | 4.9167 | 4.9609 | 4.6883 | 5.685 | 5.1667 | 7.3256 | 4.1783 | 4.7673 | 5.1433 | 4.5233 | 4 | 5.2667 | 4.3333 | 8.2512
| Median |  | 4.55 | 4.695 | 4.575 | 4.885 | 4.84 | 5.015 | 4.94 | 5.005 | 4.7 | 4.9 | 4.9492 | 4.51 | 5.68 | 5 | 7.3623 | 4.23 | 4.818 | 5.15 | 4.525 | 4 | 5.25 | 4 | 8.1345
| Std Dev. |  | 0.12111 | 0.10629 | 0.14652 | 0.19398 | 0.14976 | 0.18907 | 0.37659 | 0.3402 | 0.083666 | 0.1169 | 0.05382 | 0.5127 | 0.14598 | 0.40825 | 0.51064 | 0.16461 | 0.2897 | 0.097297 | 0.14542 | 0 | 0.13663 | 0.5164 | 0.43172
| Rel.Std.Dev. |  | 2.671% | 2.251% | 3.19% | 3.955% | 3.058% | 3.799% | 7.514% | 6.668% | 1.761% | 2.378% | 1.085% | 10.94% | 2.568% | 7.902% | 6.971% | 3.94% | 6.077% | 1.892% | 3.215% | 0% | 2.594% | 11.92% | 5.232%
| PDM3 |  | -5.879% | -1.969% | -4.633% | 1.837% | 1.664% | 3.325% | 4.052% | 5.921% | -1.381% | 2.08% | 2.998% | -2.661% | 18.03% | 7.27% | 52.09% | -13.25% | -1.021% | 6.786% | -6.087% | -16.95% | 9.346% | -10.03% | 71.31%
| Z-Score (Absolute) |  | 1.23 | 0.62 | 1.04 | 0.02 | 0.05 | 0.21 | 0.33 | 0.62 | 0.52 | 0.02 | 0.16 | 0.73 | 2.53 | NA | 7.88 | 2.39 | 0.47 | 0.76 | 1.26 | NA | 1.16 | NA | 10.9
Results from laboratories 16, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 610 (Certified Value 0.82 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.8215 | 0.818 | 0.86 | 0.79 | 0.81 | 0.76 | 0.79 | 0.81 | 0.896 | 0.8377 | 0.8116 | 0.8045 | 0.85 | 0.86 | 0.88962 | 0.84 | 0.814 | 0.84 | 0.79 | 0.78 | 1.09 | 0.81 | 0.8242 | 1
| 1 | 2 | 0.8068 | 0.809 | 0.86 | 0.81 | 0.81 | 0.8 | 0.78 | 0.84 | 0.911 | 0.8241 | 0.80941 | 0.8166 | 0.87 | 0.85 | 0.96668 | 0.83 | 0.796 | 0.85 | 0.79 | 0.79 | 1.06 | 0.8 | 0.8454 | #N/A
| 1 | 3 | 0.8206 | 0.784 | 0.82 | 0.8 | 0.82 | 0.8 | 0.78 | 0.8 | 0.905 | 0.8211 | 0.81456 | 0.8209 | 0.81 | 0.88 | 0.96502 | 0.83 | 0.827 | 0.86 | 0.77 | 0.83 | 1.11 | 0.8 | 0.8766 | 16
| 1 | 4 | 0.8094 | 0.813 | 0.83 | 0.81 | 0.8 | 0.81 | 0.79 | 0.81 | 0.901 | 0.8448 | 0.80738 | 0.82 | 0.84 | 0.85 | 0.94802 | 0.84 | 0.81 | 0.85 | 0.78 | 0.81 | 1.09 | 0.8 | 0.8378 | 0.82035
| 1 | 5 | 0.8304 | 0.847 | 0.86 | 0.8 | 0.81 | 0.81 | 0.77 | 0.8 | 0.9 | 0.8427 | 0.81388 | 0.8712 | 0.87 | 0.87 | 0.90018 | 0.83 | 0.805 | 0.85 | 0.78 | 0.78 | 1.09 | 0.8 | 0.839 | 40
| 1 | 6 | 0.8122 | 0.864 | 0.83 | 0.78 | 0.8 | 0.8 | 0.78 | 0.84 | 0.897 | 0.8256 | 0.81082 | 0.8609 | 0.82 | 0.86 | 0.88569 | 0.86 | 0.819 | 0.87 | 0.79 | 0.79 | 1.08 | 0.8 | 0.8702
| Mean |  | 0.81682 | 0.8225 | 0.84333 | 0.79833 | 0.80833 | 0.79667 | 0.78167 | 0.81667 | 0.90167 | 0.83267 | 0.81127 | 0.83235 | 0.84333 | 0.86167 | 0.92587 | 0.83833 | 0.81183 | 0.85333 | 0.78333 | 0.79667 | 1.0867 | 0.80167 | 0.84887
| Median |  | 0.8164 | 0.8155 | 0.845 | 0.8 | 0.81 | 0.8 | 0.78 | 0.81 | 0.9005 | 0.83165 | 0.81121 | 0.82045 | 0.845 | 0.86 | 0.9241 | 0.835 | 0.812 | 0.85 | 0.785 | 0.79 | 1.09 | 0.8 | 0.8422
| Std Dev. |  | 0.0089152 | 0.028627 | 0.018619 | 0.01169 | 0.0075277 | 0.018619 | 0.0075277 | 0.018619 | 0.0055737 | 0.010299 | 0.0027028 | 0.026951 | 0.025033 | 0.01169 | 0.038148 | 0.01169 | 0.010834 | 0.010328 | 0.008165 | 0.019664 | 0.01633 | 0.0040825 | 0.02032
| Rel.Std.Dev. |  | 1.091% | 3.48% | 2.208% | 1.464% | 0.9313% | 2.337% | 0.963% | 2.28% | 0.6182% | 1.237% | 0.3332% | 3.238% | 2.968% | 1.357% | 4.12% | 1.394% | 1.334% | 1.21% | 1.042% | 2.468% | 1.503% | 0.5092% | 2.394%
| PDM3 |  | -0.4305% | 0.2622% | 2.802% | -2.684% | -1.465% | -2.887% | -4.715% | -0.4488% | 9.913% | 1.502% | -1.106% | 1.463% | 2.802% | 5.037% | 12.86% | 2.192% | -1.038% | 4.021% | -4.512% | -2.887% | 32.46% | -2.277% | 3.476%
| Z-Score (Absolute) |  | 0.18 | 0 | 0.67 | 0.78 | 0.46 | 0.84 | 1.32 | 0.19 | 2.56 | 0.33 | 0.36 | 0.32 | 0.67 | 1.27 | 3.35 | 0.51 | 0.35 | 1 | 1.27 | 0.84 | 8.55 | 0.67 | 0.85
| Analytical results for Nb in OREAS 610 (Certified Value 8.27 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 8 | 8.2 | 7.5 | 8.6 | 8.2 | 7.9 | 8.9 | 9.4 | 8 | 8.62 | 8.0085 | 8.11 | 10.1 | 8.5 | 7.176 | 8.3 | 8.1 | 8.4 | 9.3 | 7 |  |  |  | 1
| 1 | 2 | 8 | 8.28 | 7.6 | 8.7 | 8.1 | 8.3 | 9 | 10 | 8.1 | 8.28 | 8.0309 | 8.07 | 10.1 | 8.5 | 7.2725 | 8.1 | 8.1 | 8.2 | 9.2 | 7 |  |  |  | 34
| 1 | 3 | 7.8 | 8.36 | 7.9 | 9 | 8.4 | 8 | 8.8 | 9.5 | 8.1 | 8.23 | 8.2096 | 8.31 | 9.7 | 8.5 | 7.2769 | 8.1 | 8.1 | 8.5 | 9 | 6 |  |  |  | 16
| 1 | 4 | 7.9 | 8.47 | 7.1 | 8.5 | 8.3 | 8.1 | 8.8 | 9.6 | 8.6 | 8.36 | 8.1954 | 7.87 | 10.4 | 9 | 7.3871 | 8.3 | 7.8 | 8.4 | 9.2 | 6 |  |  |  | 8.2736
| 1 | 5 | 8 | 8.33 | 6.5 | 8.7 | 8.1 | 8.2 | 8.7 | 9.3 | 8.3 | 8.4 | 8.2408 | 8.45 | 10 | 8.5 | 7.2755 | 8.2 | 7.8 | 8 | 9.1 | 6 |  |  |  | 41
| 1 | 6 | 7.7 | 8.12 | 6 | 8.6 | 8.2 | 8.1 | 8.5 | 9.7 | 8 | 8.36 | 8.2545 | 8.83 | 9.9 | 9 | 7.515 | 8.2 | 7.9 | 8.2 | 9.1 | 6
| Mean |  | 7.9 | 8.2933 | 7.1 | 8.6833 | 8.2167 | 8.1 | 8.7833 | 9.5833 | 8.1833 | 8.375 | 8.1566 | 8.2733 | 10.0333 | 8.6667 | 7.3172 | 8.2 | 7.9667 | 8.2833 | 9.15 | 6.3333
| Median |  | 7.95 | 8.305 | 7.3 | 8.65 | 8.2 | 8.1 | 8.8 | 9.55 | 8.1 | 8.36 | 8.2025 | 8.21 | 10.05 | 8.5 | 7.2762 | 8.2 | 8 | 8.3 | 9.15 | 6
| Std Dev. |  | 0.12649 | 0.12323 | 0.72388 | 0.17224 | 0.1169 | 0.14142 | 0.17224 | 0.24833 | 0.23166 | 0.13502 | 0.10838 | 0.33857 | 0.23381 | 0.2582 | 0.11775 | 0.089443 | 0.15055 | 0.18348 | 0.10488 | 0.5164
| Rel.Std.Dev. |  | 1.601% | 1.486% | 10.2% | 1.984% | 1.423% | 1.746% | 1.961% | 2.591% | 2.831% | 1.612% | 1.329% | 4.092% | 2.33% | 2.979% | 1.609% | 1.091% | 1.89% | 2.215% | 1.146% | 8.154%
| PDM3 |  | -4.515% | 0.2391% | -14.18% | 4.953% | -0.6876% | -2.098% | 6.162% | 15.83% | -1.09% | 1.226% | -1.413% | -0.00268% | 21.27% | 4.751% | -11.56% | -0.889% | -3.709% | 0.1182% | 10.59% | -23.45%
| Z-Score (Absolute) |  | 0.82 | 0.04 | 2.58 | 0.9 | 0.12 | 0.38 | 1.12 | 2.88 | 0.2 | 0.22 | 0.26 | 0 | 3.87 | 0.86 | 2.1 | 0.16 | 0.67 | 0.02 | 1.93 | NA
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 610 (Certified Value 18.9 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 |  |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS |  |  |  | 1
2
| 1 | 1 | 2.02: 18.4; 2.03: 20.3; 2.04: 20.21; 2.05: 15.1; 2.06: 20.8 |  |  |  | 1
| 1 | 2 | 2.02: 17.6; 2.03: 17.7; 2.04: 20.8; 2.05: 18.7; 2.06: 19.8 |  |  |  | 10
| 1 | 3 | 2.02: 18.8; 2.03: 17.6; 2.04: 19.85; 2.05: 15.8; 2.06: 20.6 |  |  |  | 16
| 1 | 4 | 2.02: 17.8; 2.03: 18.6; 2.04: 20.55; 2.05: 17.3; 2.06: 20 |  |  |  | 18.881
| 1 | 5 | 2.02: 18.3; 2.03: 19.4; 2.04: 20.15; 2.05: 17; 2.06: 20 |  |  |  | 42
| 1 | 6 | 2.02: 17.9; 2.03: 18.8; 2.04: 20.97; 2.05: 17.1; 2.06: 20.5
| Mean |  | 2.02: 18.1333; 2.03: 18.7333; 2.04: 20.4217; 2.05: 16.8333; 2.06: 20.2833
| Median |  | 2.02: 18.1; 2.03: 18.7; 2.04: 20.38; 2.05: 17.05; 2.06: 20.25
| Std Dev. |  | 2.02: 0.44572; 2.03: 1.027; 2.04: 0.4256; 2.05: 1.2549; 2.06: 0.40208
| Rel.Std.Dev. |  | 2.02: 2.458%; 2.03: 5.482%; 2.04: 2.084%; 2.05: 7.455%; 2.06: 1.982%
| PDM3 |  | 2.02: -3.96%; 2.03: -0.7821%; 2.04: 8.16%; 2.05: -10.85%; 2.06: 7.427%
| Z-Score (Absolute) |  | 2.02: 0.26; 2.03: 0; 2.04: 0.73; 2.05: 0.83; 2.06: 0.67
| Analytical results for Ni in OREAS 610 (Certified Value 24.6 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 25 | 24 | 24.2 | 22.9 | 24.8 | 23.5 | 24.9 | 24.8 | 29 | 25.1 | 25.1318 | 23.6 | 23.8 | 30 | 22.9957 | 24.2 | 24.434 | 24.6 | 24.5 | 23 | 25.7 | 27 | 29.32 | 1
| 1 | 2 | 24 | 25.2 | 25.8 | 23 | 25.1 | 25.7 | 25 | 25.4 | 29 | 24.6 | 25.0192 | 22.1 | 23 | 30 | 21.7122 | 22.6 | 24.503 | 25.1 | 25.7 | 23 | 25.9 | 27 | 26.33 | 21
| 1 | 3 | 25 | 25 | 26.1 | 23.9 | 25.7 | 23.7 | 24.7 | 23.9 | 29 | 24.8 | 25.4102 | 23.8 | 24.9 | 34 | 23.6345 | 23.9 | 26.597 | 25.4 | 25.3 | 24 | 25.1 | 27 | 25.69 | 16
| 1 | 4 | 25 | 24.7 | 25 | 22.4 | 25.7 | 23.5 | 24.6 | 24.8 | 30 | 24.5 | 24.5515 | 21.4 | 24 | 28 | 23.0497 | 24.2 | 24.204 | 25.4 | 23.8 | 23 | 25.7 | 27 | 25.63 | 24.6174
| 1 | 5 | 24 | 25.1 | 27 | 23.1 | 24.5 | 23.6 | 24.3 | 23.7 | 28 | 24.9 | 24.2863 | 26.2 | 23.9 | 32 | 25.0719 | 24.2 | 25.288 | 24.6 | 25.2 | 22 | 28.7 | 27 | 25.36 | 43
| 1 | 6 | 25 | 25.4 | 25 | 23.2 | 25.1 | 23.2 | 23.6 | 24.8 | 29 | 24.8 | 24.519 | 26.2 | 24.2 | 28 | 24.4312 | 23.1 | 25.46 | 24.6 | 25 | 23 | 25.7 | 27 | 25.4
| Mean |  | 24.6667 | 24.9 | 25.5167 | 23.0833 | 25.15 | 23.8667 | 24.5167 | 24.5667 | 29 | 24.7833 | 24.8197 | 23.8833 | 23.9667 | 30.3333 | 23.4825 | 23.7 | 25.081 | 24.95 | 24.9167 | 23 | 26.1333 | 27 | 26.2883
| Median |  | 25 | 25.05 | 25.4 | 23.05 | 25.1 | 23.55 | 24.65 | 24.8 | 29 | 24.8 | 24.7854 | 23.7 | 23.95 | 30 | 23.3421 | 24.05 | 24.8955 | 24.85 | 25.1 | 23 | 25.7 | 27 | 25.66
| Std Dev. |  | 0.5164 | 0.498 | 0.98877 | 0.48751 | 0.48062 | 0.9136 | 0.51153 | 0.64083 | 0.63246 | 0.2137 | 0.4319 | 2.0084 | 0.61536 | 2.3381 | 1.1839 | 0.68702 | 0.89516 | 0.39875 | 0.67355 | 0.63246 | 1.2863 | 0 | 1.5255
| Rel.Std.Dev. |  | 2.094% | 2% | 3.875% | 2.112% | 1.911% | 3.828% | 2.086% | 2.609% | 2.181% | 0.8623% | 1.74% | 8.409% | 2.568% | 7.708% | 5.041% | 2.899% | 3.569% | 1.598% | 2.703% | 2.75% | 4.922% | 0% | 5.803%
| PDM3 |  | 0.2002% | 1.148% | 3.653% | -6.232% | 2.164% | -3.05% | -0.4092% | -0.2061% | 17.8% | 0.6741% | 0.8218% | -2.982% | -2.643% | 23.22% | -4.61% | -3.727% | 1.883% | 1.351% | 1.216% | -6.57% | 6.158% | 9.679% | 6.788%
| Z-Score (Absolute) |  | 0.12 | 0.06 | 0.55 | 1.37 | 0.26 | 0.75 | 0.24 | 0.2 | 3.3 | 0.03 | 0 | 0.74 | 0.67 | 4.36 | 1.06 | 0.89 | 0.21 | 0.1 | 0.08 | 1.44 | 1.04 | 1.72 | 1.16
| Analytical results for P in OREAS 610 (Certified Value 554 ppm) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 534 | 520 | 550 | 510 | 530 | 510 | 530 | 550 | 570 | 574 | 492.5546 | 547 | 541 | 550 | 547.1758 | 600 | 570 | 699 | 600 | 800 | 650 | 660 | 494.3 | 1
| 1 | 2 | 527 | 520 | 540 | 570 | 520 | 530 | 530 | 550 | 590 | 567 | 531.398 | 563 | 538 | 500 | 556.8847 | 600 | 610 | 705 | 600 | 700 | 670 | 660 | 523.3 | #N/A
| 1 | 3 | 555 | 520 | 530 | 540 | 530 | 530 | 520 | 520 | 590 | 561 | 503.9493 | 564 | 560 | 550 | 552.0995 | 600 | 640 | 698 | 600 | 800 | 640 | 650 | 505.23 | 16
| 1 | 4 | 546 | 510 | 550 | 530 | 530 | 530 | 530 | 540 | 590 | 565 | 493.3463 | 554 | 553 | 550 | 555.6241 | 700 | 600 | 675 | 600 | 800 | 630 | 660 | 501.3 | 554.2234
| 1 | 5 | 547 | 530 | 550 | 530 | 530 | 540 | 520 | 540 | 580 | 566 | 529.6886 | 576 | 547 | 550 | 551.3339 | 600 | 590 | 683 | 600 | 700 | 660 | 670 | 510.3 | 44
| 1 | 6 | 539 | 540 | 550 | 520 | 520 | 540 | 510 | 560 | 580 | 555 | 494.7593 | 565 | 537 | 550 | 553.2673 | 700 | 590 | 691 | 600 | 700 | 620 | 660 | 509.3
| Mean |  | 541.3333 | 523.3333 | 545 | 533.3333 | 526.6667 | 530 | 523.3333 | 543.3333 | 583.3333 | 564.6667 | 507.616 | 561.5 | 546 | 541.6667 | 552.7309 | 633.3333 | 600 | 691.8333 | 600 | 750 | 645 | 660 | 507.2883
| Median |  | 542.5 | 520 | 550 | 530 | 530 | 530 | 525 | 545 | 585 | 565.5 | 499.3543 | 563.5 | 544 | 550 | 552.6834 | 600 | 595 | 694.5 | 600 | 750 | 645 | 660 | 507.265
| Std Dev. |  | 10.0532 | 10.328 | 8.3666 | 20.6559 | 5.164 | 10.9545 | 8.165 | 13.6626 | 8.165 | 6.3456 | 18.2318 | 9.975 | 9.1214 | 20.4124 | 3.4387 | 51.6398 | 23.6643 | 11.1788 | 0 | 54.7723 | 18.7083 | 6.3246 | 9.7804
| Rel.Std.Dev. |  | 1.857% | 1.973% | 1.535% | 3.873% | 0.9805% | 2.067% | 1.56% | 2.515% | 1.4% | 1.124% | 3.592% | 1.776% | 1.671% | 3.768% | 0.6221% | 8.154% | 3.944% | 1.616% | 0% | 7.303% | 2.901% | 0.9583% | 1.928%
| PDM3 |  | -2.326% | -5.574% | -1.664% | -3.769% | -4.972% | -4.371% | -5.574% | -1.965% | 5.252% | 1.884% | -8.409% | 1.313% | -1.484% | -2.266% | -0.2693% | 14.27% | 8.26% | 24.83% | 8.26% | 35.32% | 16.38% | 19.09% | -8.469%
| Z-Score (Absolute) |  | 0.14 | 0.67 | 0.03 | 0.38 | 0.58 | 0.48 | 0.67 | 0.08 | 1.11 | 0.56 | 1.14 | 0.46 | 0 | 0.13 | 0.2 | 2.6 | 1.61 | 4.34 | 1.61 | 6.07 | 2.95 | 3.39 | 1.15
| Analytical results for Pb in OREAS 610 (Certified Value 662 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 685 | 637.63 | 619 | 661 | 686 | 634 | 697 | 703 | 692 | 695.8 | 654.0603 | 726.22 | 689 | 696 | 666.5834 | 623.2 | 635.241 | 585.8 | 581 | 602 | 667.9 | 620 | 708.3 | 1
| 1 | 2 | 664 | 625.33 | 656 | 677 | 687 | 678 | 684 | 728 | 701 | 691.4 | 644.3905 | 705.43 | 679 | 694 | 664.4462 | 652.4 | 644.542 | 572.9 | 563 | 592 | 643.4 | 600 | 709.6 | 23
| 1 | 3 | 666 | 628.1 | 640 | 669 | 697 | 674 | 687 | 689 | 722 | 677.7 | 661.9919 | 715.42 | 692 | 701 | 665.7268 | 636.7 | 635.145 | 576 | 588 | 598 | 647.7 | 620 | 710.9 | 16
| 1 | 4 | 674 | 640.47 | 629 | 673 | 687 | 680 | 692 | 700 | 729 | 688.9 | 647.5708 | 695.04 | 682 | 723 | 666.3761 | 644.9 | 646.849 | 593 | 594 | 591 | 629.4 | 610 | 712.2 | 661.6457
| 1 | 5 | 685 | 648.78 | 638 | 664 | 682 | 676 | 680 | 698 | 706 | 685.5 | 655.2061 | 767 | 678 | 735 | 672.0404 | 645.9 | 645.365 | 599.1 | 565 | 584 | 627.7 | 610 | 712.5 | 45
| 1 | 6 | 681 | 666.41 | 619 | 657 | 685 | 672 | 681 | 726 | 722 | 694 | 656.2066 | 733 | 673 | 723 | 665.71 | 630.5 | 641.327 | 581.9 | 587 | 595 | 634.5 | 620 | 710.8
| Mean |  | 675.8333 | 641.12 | 633.5 | 666.8333 | 687.3333 | 669 | 686.8333 | 707.3333 | 712 | 688.8833 | 653.2377 | 723.685 | 682.1667 | 712 | 666.8138 | 638.9333 | 641.4115 | 584.7833 | 579.6667 | 593.6667 | 641.7667 | 613.3333 | 710.7167
| Median |  | 677.5 | 639.05 | 633.5 | 666.5 | 686.5 | 675 | 685.5 | 701.5 | 714 | 690.15 | 654.6332 | 720.82 | 680.5 | 712 | 666.0514 | 640.8 | 642.9345 | 583.85 | 584 | 593.5 | 638.95 | 615 | 710.85
| Std Dev. |  | 9.3256 | 15.0313 | 14.2092 | 7.5476 | 5.0859 | 17.3781 | 6.6156 | 15.9457 | 14.4637 | 6.5859 | 6.3307 | 25.2645 | 7.1391 | 17.1581 | 2.6671 | 10.8548 | 5.1449 | 10.0019 | 12.8323 | 6.2183 | 14.9962 | 8.165 | 1.5817
| Rel.Std.Dev. |  | 1.38% | 2.345% | 2.243% | 1.132% | 0.74% | 2.598% | 0.9632% | 2.254% | 2.031% | 0.956% | 0.9691% | 3.491% | 1.047% | 2.41% | 0.4% | 1.699% | 0.8021% | 1.71% | 2.214% | 1.047% | 2.337% | 1.331% | 0.2225%
| PDM3 |  | 2.144% | -3.102% | -4.254% | 0.784% | 3.882% | 1.112% | 3.807% | 6.905% | 7.61% | 4.117% | -1.271% | 9.377% | 3.102% | 7.61% | 0.7811% | -3.433% | -3.058% | -11.62% | -12.39% | -10.27% | -3.004% | -7.302% | 7.417%
| Z-Score (Absolute) |  | 0.24 | 0.67 | 0.87 | 0 | 0.54 | 0.06 | 0.52 | 1.06 | 1.18 | 0.58 | 0.36 | 1.49 | 0.4 | 1.18 | 0 | 0.73 | 0.67 | 2.15 | 2.29 | 1.92 | 0.66 | 1.4 | 1.15
| Analytical results for Pr in OREAS 610 (Certified Value 5.09 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
2
| 1 | 1 | 2.02: 4.8; 2.03: 4.9; 2.04: 5.69; 2.05: 4.1; 2.06: 6 | 1
| 1 | 2 | 2.02: 5; 2.03: 4.3; 2.04: 5.77; 2.05: 4.7; 2.06: 5.7 | 12
| 1 | 3 | 2.02: 5.1; 2.03: 4.2; 2.04: 5.68; 2.05: 4.3; 2.06: 5.84 | 16
| 1 | 4 | 2.02: 4.9; 2.03: 4.6; 2.04: 5.76; 2.05: 4.6; 2.06: 5.76 | 5.0923
| 1 | 5 | 2.02: 5.1; 2.03: 4.8; 2.04: 5.6; 2.05: 4.3; 2.06: 5.66 | 46
| 1 | 6 | 2.02: 4.9; 2.03: 4.8; 2.04: 5.77; 2.05: 4.3; 2.06: 5.84
| Mean |  | 2.02: 4.9667; 2.03: 4.6; 2.04: 5.7117; 2.05: 4.3833; 2.06: 5.8
| Median |  | 2.02: 4.95; 2.03: 4.7; 2.04: 5.725; 2.05: 4.3; 2.06: 5.8
| Std Dev. |  | 2.02: 0.12111; 2.03: 0.28983; 2.04: 0.067946; 2.05: 0.22286; 2.06: 0.12198
| Rel.Std.Dev. |  | 2.02: 2.438%; 2.03: 6.301%; 2.04: 1.19%; 2.05: 5.084%; 2.06: 2.103%
| PDM3 |  | 2.02: -2.468%; 2.03: -9.668%; 2.04: 12.16%; 2.05: -13.92%; 2.06: 13.9%
| Z-Score (Absolute) |  | 2.02: 0; 2.03: 0.42; 2.04: 0.86; 2.05: 0.67; 2.06: 0.96
| Analytical results for Rb in OREAS 610 (Certified Value 65 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 67.2 | 66.4 | 61.2 | 64.2 | 62.5 | 62.7 | 65.7 | 65.8 | 62 | 67.87 | 69.2352 | 67.7 | 60.2 | 68.4 | 60.4 | 60.5 | 63.6 | 73.6 | [5 empty] | 1
| 1 | 2 | 66.2 | 65.1 | 61.3 | 63.5 | 62.9 | 69.6 | 66 | 68.8 | 63.9 | 67.35 | 69.1811 | 67.2 | 60.6 | 71.2 | 61.9 | 64.6 | 63.1 | 72.7 | [5 empty] | 39
| 1 | 3 | 66.2 | 64 | 60.8 | 65.5 | 66.4 | 67.5 | 65 | 66.1 | 62.3 | 65.16 | 69.8519 | 66.8 | 61.3 | 70.8 | 60.3 | 61.3 | 66.1 | 72.2 | [5 empty] | 16
| 1 | 4 | 66.6 | 66 | 59.9 | 62.8 | 62.9 | 68.2 | 65.3 | 67.8 | 63.1 | 67.08 | 70.308 | 68.2 | 61.7 | 70.6 | 62.9 | 59.8 | 63 | 72.1 | [5 empty] | 65.2767
| 1 | 5 | 66.7 | 64.5 | 59.5 | 64.5 | 62.1 | 70.2 | 63.4 | 64.2 | 63.8 | 66.84 | 68.8347 | 80.9 | 60.6 | 70.4 | 62.9 | 60.9 | 62.4 | 71 | [5 empty] | 47
| 1 | 6 | 66.9 | 66.3 | 57.5 | 64.8 | 63.8 | 67.1 | 61.4 | 67.9 | 63.7 | 67.18 | 69.8668 | 79.3 | 61.6 | 71.2 | 61.7 | 59.8 | 63.6 | 71.6
| Mean |  | 66.6333 | 65.3833 | 60.0333 | 64.2167 | 63.4333 | 67.55 | 64.4667 | 66.7667 | 63.1333 | 66.9133 | 69.5463 | 71.6833 | 61 | 70.4333 | 61.6833 | 61.15 | 63.6333 | 72.2
| Median |  | 66.65 | 65.55 | 60.35 | 64.35 | 62.9 | 67.85 | 65.15 | 66.95 | 63.4 | 67.13 | 69.5436 | 67.95 | 60.95 | 70.7 | 61.8 | 60.7 | 63.35 | 72.15
| Std Dev. |  | 0.39328 | 1.0028 | 1.4334 | 0.95795 | 1.5591 | 2.6599 | 1.7546 | 1.7002 | 0.8165 | 0.92567 | 0.55005 | 6.556 | 0.61644 | 1.0463 | 1.1462 | 1.7919 | 1.2879 | 0.89666
| Rel.Std.Dev. |  | 0.5902% | 1.534% | 2.388% | 1.492% | 2.458% | 3.938% | 2.722% | 2.546% | 1.293% | 1.383% | 0.7909% | 9.146% | 1.011% | 1.485% | 1.858% | 2.93% | 2.024% | 1.242%
| PDM3 |  | 2.078% | 0.1633% | -8.033% | -1.624% | -2.824% | 3.482% | -1.241% | 2.282% | -3.284% | 2.507% | 6.541% | 9.815% | -6.552% | 7.9% | -5.505% | -6.322% | -2.518% | 10.61%
| Z-Score (Absolute) |  | 0.5 | 0.13 | 1.43 | 0.21 | 0.44 | 0.77 | 0.13 | 0.54 | 0.52 | 0.58 | 1.35 | 1.98 | 1.15 | 1.61 | 0.95 | 1.1 | 0.38 | 2.13
| Analytical results for Re in OREAS 610 (Indicative Value < 2 ppb) | [24 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.23 | [7 empty] | ppb
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | < 50 | 5 | 2 | < 2 | < 2 | < 2 | < 2 | < 2 | < 50 | 4 | < 50 | < 2 | 2 | < 100 | < 2 | 8 | [7 empty] | 1
| 1 | 2 | < 50 | < 2 | 2 | < 2 | < 2 | < 2 | < 2 | < 2 | < 50 | 4 | < 50 | < 2 | 2 | < 100 | < 2 | < 5 | [7 empty] | 14
| 1 | 3 | < 50 | 4 | 1 | < 2 | 2 | < 2 | < 2 | < 2 | < 50 | < 2 | < 50 | < 2 | 2 | < 100 | < 2 | < 5 | [7 empty] | 16
| 1 | 4 | < 50 | 4 | 2 | 3 | 2 | < 2 | 2 | < 2 | < 50 | < 2 | < 50 | < 2 | 3 | < 100 | < 2 | < 5 | [7 empty] | < 2
| 1 | 5 | < 50 | < 2 | 2 | < 2 | 2 | < 2 | < 2 | < 2 | < 50 | < 2 | < 50 | < 2 | 2 | < 100 | < 2 | < 5 | [7 empty] | 11
| 1 | 6 | < 50 | 3 | 2 | 3 | < 2 | < 2 | < 2 | < 2 | < 50 | 2 | < 50 | < 2 | 3 | < 100 | < 2 | 6
| Mean |  | 2.03: 4; 2.04: 1.8333; 2.05: 3; 2.06: 2; 2.08: 2; 2.11: 3.3333; 2.15: 2.3333; 2.23: 7
| Median |  | 2.03: 4; 2.04: 2; 2.05: 3; 2.06: 2; 2.08: 2; 2.11: 4; 2.15: 2; 2.23: 7
| Std Dev. |  | 2.03: 0.8165; 2.04: 0.40825; 2.05: 0; 2.06: 0; 2.11: 1.1547; 2.15: 0.5164; 2.23: 1.4142
| Rel.Std.Dev. |  | 2.03: 20.41%; 2.04: 22.27%; 2.05: 0%; 2.06: 0%; 2.11: 34.64%; 2.15: 22.13%; 2.23: 20.2%
PDM3
| Z-Score (Absolute) |  | 15.51 | 0.67 | 0.11 | 0.22 | 0.34 | 0.67 | 0.56 | 0.67 | 15.51 | 0.11 | 15.51 | 0.67 | 0.22 | 32.37 | 0.67 | 1.35
| Analytical results for S in OREAS 610 (Certified Value 4.07 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 4.0673 | 3.82 | 4.17 | 4.17 | 4.13 | 3.91 | 4.02 | 4.22 | 4.01 | 4.0728 | 4.1656 | 4.1912 | 3.82 | 4.03 | 3.9912 | 3.95 | 3.991 | 4.02 | 3.96 | 4.36 | 3.92 | 4.2 | 3.967 | 1
| 1 | 2 | 3.9502 | 3.84 | 4.12 | 4.25 | 4.1 | 4.1 | 3.98 | 4.36 | 4.02 | 3.9938 | 4.2076 | 4.1966 | 3.83 | 4.15 | 4.0055 | 3.88 | 3.971 | 4.11 | 4.15 | 4.27 | 3.97 | 4.2 | 4.036 | 16
| 1 | 3 | 4.0572 | 3.93 | 4.05 | 4.23 | 4.13 | 4.09 | 3.96 | 4.16 | 4.06 | 4.0391 | 4.2247 | 4.2477 | 3.76 | 4.22 | 4.0189 | 3.85 | 4.237 | 4.03 | 3.94 | 4.37 | 3.94 | 4.2 | 4.105 | 16
| 1 | 4 | 4.0177 | 3.77 | 4.23 | 4.28 | 4.14 | 4.16 | 3.99 | 4.22 | 4.05 | 4.0941 | 4.2597 | 4.2575 | 3.75 | 4.16 | 4.0534 | 3.93 | 3.939 | 4.01 | 3.75 | 4.31 | 3.95 | 4.2 | 4.174 | 4.0693
| 1 | 5 | 4.0881 | 3.76 | 4.2 | 4.21 | 4.09 | 4.16 | 3.93 | 4.19 | 3.99 | 4.1123 | 4.17 | 4.3663 | 3.82 | 4.2 | 4.0255 | 3.9 | 3.943 | 4.07 | 3.93 | 4.23 | 3.99 | 4.19 | 4.043 | 48
| 1 | 6 | 4.005 | 3.89 | 4.13 | 4.16 | 4.08 | 4.11 | 3.97 | 4.39 | 3.87 | 4.0699 | 4.2207 | 4.2787 | 3.74 | 4.16 | 3.9748 | 3.97 | 3.996 | 4 | 3.84 | 4.26 | 4 | 4.19 | 4.012
| Mean |  | 4.0309 | 3.835 | 4.15 | 4.2167 | 4.1117 | 4.0883 | 3.975 | 4.2567 | 4 | 4.0637 | 4.2081 | 4.2563 | 3.7867 | 4.1533 | 4.0115 | 3.9133 | 4.0128 | 4.04 | 3.9283 | 4.3 | 3.9617 | 4.1967 | 4.0562
| Median |  | 4.0374 | 3.83 | 4.15 | 4.22 | 4.115 | 4.105 | 3.975 | 4.22 | 4.015 | 4.0713 | 4.2141 | 4.2526 | 3.79 | 4.16 | 4.0122 | 3.915 | 3.981 | 4.025 | 3.935 | 4.29 | 3.96 | 4.2 | 4.0395
| Std Dev. |  | 0.050291 | 0.066558 | 0.064187 | 0.046332 | 0.024833 | 0.092394 | 0.030166 | 0.094798 | 0.068702 | 0.042187 | 0.035657 | 0.063999 | 0.040825 | 0.066232 | 0.027592 | 0.045019 | 0.11233 | 0.041952 | 0.13408 | 0.056569 | 0.030605 | 0.005164 | 0.073145
| Rel.Std.Dev. |  | 1.248% | 1.736% | 1.547% | 1.099% | 0.604% | 2.26% | 0.7589% | 2.227% | 1.718% | 1.038% | 0.8473% | 1.504% | 1.078% | 1.595% | 0.6878% | 1.15% | 2.799% | 1.038% | 3.413% | 1.316% | 0.7725% | 0.123% | 1.803%
| PDM3 |  | -0.9437% | -5.758% | 1.983% | 3.621% | 1.041% | 0.4672% | -2.318% | 4.604% | -1.703% | -0.1389% | 3.409% | 4.596% | -6.946% | 2.065% | -1.42% | -3.833% | -1.388% | -0.7205% | -3.465% | 5.669% | -2.645% | 3.129% | -0.3232%
| Z-Score (Absolute) |  | 0.18 | 1.58 | 0.67 | 1.15 | 0.4 | 0.23 | 0.58 | 1.43 | 0.4 | 0.05 | 1.08 | 1.43 | 1.92 | 0.69 | 0.32 | 1.02 | 0.31 | 0.12 | 0.91 | 1.74 | 0.67 | 1 | 0
| Analytical results for Sb in OREAS 610 (Certified Value 299 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 267.4 | 306.01 | 37.4 | 306 | 295 | 289 | 305 | 313 | 302 | 317.36 | 310.9675 | 315.33 | 331 | 342 | 344.6148 | 279.2 | 306.983 | 293.3 | 279 | 299 | 262.5 | 270 | 295.2 | 1
| 1 | 2 | 264.8 | 302.99 | 54.6 | 313 | 289 | 305 | 298 | 320 | 302 | 309.43 | 306.1842 | 321.15 | 326 | 354 | 332.1157 | 257.5 | 291.724 | 298.9 | 283 | 293 | 261.2 | 280 | 295.4 | 41
| 1 | 3 | 265.2 | 305.24 | 82.3 | 313 | 297 | 305 | 300 | 306 | 296 | 311.43 | 293.1202 | 323.83 | 320 | 359 | 352.4893 | 278.4 | 311.701 | 294.7 | 287 | 298 | 258.1 | 270 | 299.6 | 16
| 1 | 4 | 263.4 | 313.55 | 73.7 | 319 | 298 | 314 | 297 | 308 | 299 | 305.5 | 315.041 | 319.4 | 333 | 362 | 365.5305 | 291.4 | 306.725 | 287.6 | 282 | 299 | 259.7 | 280 | 293.8 | 298.8757
| 1 | 5 | 264.8 | 309.85 | 47.4 | 309 | 295 | 309 | 295 | 305 | 301 | 314.23 | 295.5314 | 315 | 321 | 360 | 325.7565 | 277.6 | 313.477 | 290.8 | 276 | 289 | 254.2 | 280 | 298 | 49
| 1 | 6 | 263 | 309.3 | 43.5 | 309 | 292 | 303 | 294 | 325 | 303 | 311.7 | 312.4935 | 309.5 | 326 | 363 | 345.1689 | 266.9 | 298.981 | 294.6 | 279 | 292 | 256 | 270 | 359.2
| Mean |  | 264.7667 | 307.8233 | 56.4833 | 311.5 | 294.3333 | 304.1667 | 298.1667 | 312.8333 | 300.5 | 311.6083 | 305.5563 | 317.3683 | 326.1667 | 356.6667 | 344.2793 | 275.1667 | 304.9318 | 293.3167 | 281 | 295 | 258.6167 | 275 | 306.8667
| Median |  | 264.8 | 307.655 | 51 | 311 | 295 | 305 | 297.5 | 310.5 | 301.5 | 311.565 | 308.5758 | 317.365 | 326 | 359.5 | 344.8919 | 278 | 306.854 | 293.95 | 280.5 | 295.5 | 258.9 | 275 | 296.7
| Std Dev. |  | 1.5565 | 3.8054 | 17.786 | 4.5497 | 3.3267 | 8.4004 | 3.9707 | 8.1343 | 2.5884 | 4.051 | 9.1969 | 5.1353 | 5.1929 | 7.8401 | 14.2195 | 11.6328 | 8.1991 | 3.8405 | 3.8471 | 4.2426 | 3.148 | 5.4772 | 25.7237
| Rel.Std.Dev. |  | 0.5879% | 1.236% | 31.49% | 1.461% | 1.13% | 2.762% | 1.332% | 2.6% | 0.8614% | 1.3% | 3.01% | 1.618% | 1.592% | 2.198% | 4.13% | 4.228% | 2.689% | 1.309% | 1.369% | 1.438% | 1.217% | 1.992% | 8.383%
| PDM3 |  | -11.41% | 2.994% | -81.1% | 4.224% | -1.52% | 1.77% | -0.2372% | 4.67% | 0.5435% | 4.26% | 2.235% | 6.187% | 9.131% | 19.34% | 15.19% | -7.933% | 2.026% | -1.86% | -5.981% | -1.297% | -13.47% | -7.989% | 2.674%
| Z-Score (Absolute) |  | 2.7 | 0.25 | 16.98 | 0.5 | 0.67 | 0 | 0.41 | 0.59 | 0.25 | 0.51 | 0.1 | 0.91 | 1.51 | 3.6 | 2.75 | 1.99 | 0.05 | 0.74 | 1.59 | 0.63 | 3.12 | 2 | 0.19
| Analytical results for Sc in OREAS 610 (Certified Value 3.06 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3 | 3.5 | 4 | 3.1 | 3.1 | 3.1 | 3.1 | 3.2 | 3 | 3.5 | 2.9919 | 2.6 | 3.4 | 3 | < 5 | 2.9 | 2.899 | 2.8 | 2.8 | 2.9 | 3.3 |  |  | 1
| 1 | 2 | 3 | 3.2 | 4 | 3.3 | 3.1 | 3.3 | 3.1 | 3.3 | 3 | 3.2 | 2.9021 | 2.8 | 3.4 | 4 | < 5 | 3 | 2.89 | 2.8 | 2.8 | 2.9 | 3.1 |  |  | 42
| 1 | 3 | 3 | 3.5 | 3 | 3.3 | 3.2 | 3.2 | 3 | 3.1 | 3 | 3.2 | 3.0274 | 2.7 | 3.5 | 4 | < 5 | 2.9 | 2.84 | 2.6 | 2.9 | 2.9 | 3.5 |  |  | 16
| 1 | 4 | 3 | 3.4 | 4 | 3.1 | 3.2 | 3.3 | 3.2 | 3.1 | 3 | 3.1 | 3.1545 | 2.7 | 3.5 | 3 | < 5 | 3 | 2.739 | 2.6 | 2.8 | 2.9 | 3.1 |  |  | 3.0648
| 1 | 5 | 3 | 3.4 | 4 | 3.1 | 3.1 | 3.3 | 3 | 3 | 3 | 3.2 | 2.8871 | 3.2 | 3.5 | 4 | < 5 | 2.9 | 2.831 | 2.5 | 2.7 | 2.8 | 3.2 |  |  | 50
| 1 | 6 | 3 | 3.3 | 4 | 3.1 | 3.1 | 3.2 | 3 | 3.2 | 3 | 3.3 | 3.1689 | 3.1 | 3.4 | 3 | < 5 | 3 | 2.79 | 2.6 | 2.8 | 2.9 | 3.4
| Mean |  | 3 | 3.3833 | 3.8333 | 3.1667 | 3.1333 | 3.2333 | 3.0667 | 3.15 | 3 | 3.25 | 3.022 | 2.85 | 3.45 | 3.5 |  | 2.95 | 2.8315 | 2.65 | 2.8 | 2.8833 | 3.2667
| Median |  | 3 | 3.4 | 4 | 3.1 | 3.1 | 3.25 | 3.05 | 3.15 | 3 | 3.2 | 3.0097 | 2.75 | 3.45 | 3.5 |  | 2.95 | 2.8355 | 2.6 | 2.8 | 2.9 | 3.25
| Std Dev. |  | 0 | 0.1169 | 0.40825 | 0.10328 | 0.05164 | 0.08165 | 0.08165 | 0.10488 | 0 | 0.13784 | 0.12056 | 0.2429 | 0.054772 | 0.54772 |  | 0.054772 | 0.060547 | 0.12247 | 0.063246 | 0.040825 | 0.1633
| Rel.Std.Dev. |  | 0% | 3.455% | 10.65% | 3.261% | 1.648% | 2.525% | 2.662% | 3.33% | 0% | 4.241% | 3.99% | 8.523% | 1.588% | 15.65% |  | 1.857% | 2.138% | 4.622% | 2.259% | 1.416% | 4.999%
| PDM3 |  | -2.114% | 10.39% | 25.08% | 3.324% | 2.236% | 5.499% | 0.06088% | 2.78% | -2.114% | 6.043% | -1.397% | -7.009% | 12.57% | 14.2% |  | -3.746% | -7.612% | -13.53% | -8.64% | -5.921% | 6.587%
| Z-Score (Absolute) |  | NA | 1.16 | NA | 0.37 | 0.25 | 0.61 | 0 | 0.31 | NA | 0.67 | 0.16 | 0.8 | 1.41 | NA | 2.08 | 0.43 | 0.86 | 1.53 | 0.98 | 0.67 | 0.74
Results from laboratories 2, 4, 10 & 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 610 (Certified Value 29.1 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
1
| 1 | 1 | 28 | 28.5 | 33 | 29 | 29 | 25 | 31 | 33 | 24 | 29.6 | 30.9307 | 24.9 | 34.3 | 25 | 27.9631 | 18.4 | 29 | 24 | [5 empty] | 1
| 1 | 2 | 27 | 30.2 | 32.9 | 27 | 30 | 28 | 32 | 32 | 25 | 28.6 | 31.4481 | 25.6 | 34.3 | 25 | 31.3714 | 20 | 30 | 24 | [5 empty] | 43
| 1 | 3 | 28 | 30 | 34.3 | 30 | 30 | 28 | 31 | 32 | 24 | 27.9 | 30.1539 | 26.9 | 32.2 | 25 | 28.6161 | 19 | 32 | 25 | [5 empty] | 16
| 1 | 4 | 28 | 29.5 | 33.1 | 26 | 31 | 29 | 32 | 32 | 25 | 27.6 | 32.1499 | 25.4 | 33.6 | 30 | 29.9374 | 18.9 | 30 | 23 | [5 empty] | 29.123
| 1 | 5 | 28 | 29.5 | 33.2 | 30 | 31 | 28 | 31 | 32 | 24 | 28.3 | 30.0042 | 26.6 | 33.9 | 25 | 31.2643 | 18 | 29 | 24 | [5 empty] | 51
| 1 | 6 | 28 | 31.1 | 31.8 | 31 | 30 | 28 | 30 | 33 | 26 | 28.7 | 31.5283 | 26.9 | 33.2 | 25 | 34.3759 | 19.6 | 32 | 24
| Mean |  | 27.8333 | 29.8 | 33.05 | 28.8333 | 30.1667 | 27.6667 | 31.1667 | 32.3333 | 24.6667 | 28.45 | 31.0358 | 26.05 | 33.5833 | 25.8333 | 30.588 | 18.9833 | 30.3333 | 24
| Median |  | 28 | 29.75 | 33.05 | 29.5 | 30 | 28 | 31 | 32 | 24.5 | 28.45 | 31.1894 | 26.1 | 33.75 | 25 | 30.6008 | 18.95 | 30 | 24
| Std Dev. |  | 0.40825 | 0.86718 | 0.79687 | 1.9408 | 0.75277 | 1.3663 | 0.75277 | 0.5164 | 0.8165 | 0.70071 | 0.83747 | 0.85965 | 0.79854 | 2.0412 | 2.3074 | 0.73869 | 1.3663 | 0.63246
| Rel.Std.Dev. |  | 1.467% | 2.91% | 2.411% | 6.731% | 2.495% | 4.938% | 2.415% | 1.597% | 3.31% | 2.463% | 2.698% | 3.3% | 2.378% | 7.902% | 7.544% | 3.891% | 4.504% | 2.635%
| PDM3 |  | -4.428% | 2.325% | 13.48% | -0.9945% | 3.584% | -5.001% | 7.017% | 11.02% | -15.3% | -2.311% | 6.568% | -10.55% | 15.32% | -11.3% | 5.031% | -34.82% | 4.156% | -17.59%
| Z-Score (Absolute) |  | 0.56 | 0.18 | 1.41 | 0.18 | 0.32 | 0.62 | 0.7 | 1.14 | 1.76 | 0.33 | 0.65 | 1.23 | 1.61 | 1.32 | 0.48 | 3.9 | 0.38 | 2.01
| Analytical results for Sm in OREAS 610 (Certified Value 3.68 ppm) | [24 empty] | Certified
Sm | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | [5 empty] | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | [5 empty] | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.02: 3.5; 2.03: 3.2; 2.04: 4; 2.05: 3.1; 2.06: 3.8 | [5 empty] | 1
| 1 | 2 | 2.02: 3.6; 2.03: 3.5; 2.04: 3.98; 2.05: 3.6; 2.06: 3.8 | [5 empty] | 8
| 1 | 3 | 2.02: 3.6; 2.03: 3.1; 2.04: 3.82; 2.05: 3.3; 2.06: 4 | [5 empty] | 16
| 1 | 4 | 2.02: 3.6; 2.03: 3.7; 2.04: 4.01; 2.05: 3.4; 2.06: 3.95 | [5 empty] | 3.6759
| 1 | 5 | 2.02: 3.5; 2.03: 3.5; 2.04: 4.03; 2.05: 3.7; 2.06: 3.95 | [5 empty] | 52
| 1 | 6 | 2.02: 3.5; 2.03: 3.6; 2.04: 3.96; 2.05: 3.8; 2.06: 4
| Mean |  | 2.02: 3.55; 2.03: 3.4333; 2.04: 3.9667; 2.05: 3.4833; 2.06: 3.9167
| Median |  | 2.02: 3.55; 2.03: 3.5; 2.04: 3.99; 2.05: 3.5; 2.06: 3.95
| Std Dev. |  | 2.02: 0.054772; 2.03: 0.23381; 2.04: 0.075807; 2.05: 0.26394; 2.06: 0.093095
| Rel.Std.Dev. |  | 2.02: 1.543%; 2.03: 6.81%; 2.04: 1.911%; 2.05: 7.577%; 2.06: 2.377%
| PDM3 |  | 2.02: -3.424%; 2.03: -6.598%; 2.04: 7.911%; 2.05: -5.238%; 2.06: 6.551%
| Z-Score (Absolute) |  | 2.02: 0; 2.03: 0.67; 2.04: 2.41; 2.05: 0.39; 2.06: 2.12
| Analytical results for Sn in OREAS 610 (Certified Value 27.1 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 26.8 | 29.8 | 27 | 26 | 26.4 | 25.5 | 27.4 | 29.4 | 24.1 | 28.2 | 29.4335 | 30.4 | 28.8 | 28 | 28.1054 | 25.7 | 26.839 | 25.4 | 24.9 | 24 | 31.1 | 26 | 22.77 | 1
| 1 | 2 | 26.7 | 27.9 | 26 | 27 | 26.5 | 27.5 | 27.7 | 29.7 | 24.7 | 28.1 | 29.137 | 30.5 | 28.3 | 29 | 27.7675 | 24.3 | 26.394 | 26.8 | 25 | 23 | 30.9 | 25 | 21.24 | 18
| 1 | 3 | 26.8 | 28.6 | 26 | 27.7 | 27.7 | 26.8 | 27.4 | 25.9 | 23.2 | 27.6 | 26.9815 | 30.9 | 28.9 | 29 | 26.8519 | 25.4 | 27.847 | 26.7 | 25.3 | 24 | 30.1 | 25 | 21.71 | 16
| 1 | 4 | 27.2 | 30.5 | 25 | 26.6 | 27 | 27.2 | 28 | 30 | 24.5 | 27.4 | 27.7728 | 32.3 | 28.7 | 29 | 30.0171 | 26 | 26.228 | 25.4 | 25.4 | 23 | 30 | 26 | 23.18 | 27.1181
| 1 | 5 | 27 | 28.4 | 27 | 27.6 | 25.6 | 27.3 | 26.6 | 26.1 | 24.2 | 28 | 29.5236 | 29.9 | 28.4 | 29 | 26.3815 | 25.6 | 27.079 | 25.8 | 24.7 | 22 | 30.3 | 25 | 22.65 | 53
| 1 | 6 | 26.4 | 29.3 | 26 | 26.8 | 27 | 26.8 | 26.9 | 26.5 | 24.1 | 27.8 | 26.5614 | 31.5 | 28 | 30 | 27.5718 | 24.8 | 26.775 | 25.9 | 24.6 | 22 | 30.1 | 26 | 21.12
| Mean |  | 26.8167 | 29.0833 | 26.1667 | 26.95 | 26.7 | 26.85 | 27.3333 | 27.9333 | 24.1333 | 27.85 | 28.235 | 30.9167 | 28.5167 | 29 | 27.7825 | 25.3 | 26.8603 | 26 | 24.9833 | 23 | 30.4167 | 25.5 | 22.1117
| Median |  | 26.8 | 28.95 | 26 | 26.9 | 26.75 | 27 | 27.4 | 27.95 | 24.15 | 27.9 | 28.4549 | 30.7 | 28.55 | 29 | 27.6696 | 25.5 | 26.807 | 25.85 | 24.95 | 23 | 30.2 | 25.5 | 22.18
| Std Dev. |  | 0.27142 | 0.96626 | 0.75277 | 0.63797 | 0.70993 | 0.71764 | 0.51251 | 1.9541 | 0.5164 | 0.30822 | 1.3036 | 0.86352 | 0.34303 | 0.63246 | 1.2626 | 0.63246 | 0.57382 | 0.61644 | 0.31885 | 0.89443 | 0.46655 | 0.54772 | 0.86823
| Rel.Std.Dev. |  | 1.012% | 3.322% | 2.877% | 2.367% | 2.659% | 2.673% | 1.875% | 6.996% | 2.14% | 1.107% | 4.617% | 2.793% | 1.203% | 2.181% | 4.545% | 2.5% | 2.136% | 2.371% | 1.276% | 3.889% | 1.534% | 2.148% | 3.927%
| PDM3 |  | -1.111% | 7.247% | -3.508% | -0.6198% | -1.542% | -0.9886% | 0.7938% | 3.006% | -11.01% | 2.699% | 4.119% | 14.01% | 5.157% | 6.94% | 2.45% | -6.704% | -0.9505% | -4.123% | -7.872% | -15.19% | 12.16% | -5.967% | -18.46%
| Z-Score (Absolute) |  | 0.02 | 1.1 | 0.34 | 0.04 | 0.08 | 0.01 | 0.23 | 0.53 | 1.35 | 0.49 | 0.68 | 2.01 | 0.82 | 1.06 | 0.46 | 0.77 | 0 | 0.43 | 0.93 | 1.91 | 1.76 | 0.67 | 2.35
| Analytical results for Sr in OREAS 610 (Certified Value 306 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 312.6 | 263 | 218 | 315 | 318 | 301 | 320 | 322 | 303 | 322.27 | 333.27 | 212 | 313 | 325 | 338.6367 | 241.6 | 263.51 | 249.4 | 320 | 296 | 269 | 210 | 329.7 | 1
| 1 | 2 | 306.9 | 261 | 192 | 320 | 316 | 325 | 301 | 330 | 320 | 323.82 | 337.0781 | 227 | 324 | 326 | 343.9856 | 271.2 | 269.17 | 259.3 | 313 | 301 | 263 | 210 | 332.5 | 19
| 1 | 3 | 301.5 | 262 | 189 | 317 | 322 | 318 | 301 | 320 | 316 | 321.84 | 321.9522 | 224 | 300 | 334 | 345.0935 | 260.5 | 262.055 | 254.8 | 305 | 311 | 267 | 220 | 335.3 | 16
| 1 | 4 | 311.1 | 272 | 198 | 321 | 320 | 325 | 314 | 319 | 318 | 320.78 | 338.8006 | 233 | 305 | 343 | 343.8128 | 262.1 | 254.486 | 243.9 | 316 | 317 | 266 | 210 | 335.1 | 305.75
| 1 | 5 | 313.7 | 258 | 179 | 317 | 320 | 327 | 298 | 320 | 316 | 315.92 | 311.7629 | 312 | 318 | 344 | 345.579 | 261.4 | 257.62 | 257.1 | 320 | 297 | 273 | 210 | 333.9 | 54
| 1 | 6 | 305.5 | 261 | 183 | 314 | 316 | 322 | 309 | 334 | 322 | 323.95 | 313.5074 | 250 | 298 | 348 | 345.3212 | 259.1 | 261.122 | 258.1 | 318 | 297 | 280 | 210 | 333.7
| Mean |  | 308.55 | 262.8333 | 193.1667 | 317.3333 | 318.6667 | 319.6667 | 307.1667 | 324.1667 | 315.8333 | 321.43 | 326.0619 | 243 | 309.6667 | 336.6667 | 343.7381 | 259.3167 | 261.3272 | 253.7667 | 315.3333 | 303.1667 | 269.6667 | 211.6667 | 333.3667
| Median |  | 309 | 261.5 | 190.5 | 317 | 319 | 323.5 | 305 | 321 | 317 | 322.055 | 327.6111 | 230 | 309 | 338.5 | 344.5396 | 260.95 | 261.5885 | 255.95 | 317 | 299 | 268 | 210 | 333.8
| Std Dev. |  | 4.715 | 4.7924 | 13.8768 | 2.7325 | 2.4221 | 9.6678 | 8.6583 | 6.2743 | 6.7057 | 2.957 | 11.9545 | 36.0222 | 10.3666 | 9.7912 | 2.6009 | 9.6812 | 5.0476 | 5.9665 | 5.7155 | 8.7731 | 6.0553 | 4.0825 | 2.0656
| Rel.Std.Dev. |  | 1.528% | 1.823% | 7.184% | 0.8611% | 0.7601% | 3.024% | 2.819% | 1.936% | 2.123% | 0.9199% | 3.666% | 14.82% | 3.348% | 2.908% | 0.7567% | 3.733% | 1.932% | 2.351% | 1.813% | 2.894% | 2.245% | 1.929% | 0.6196%
| PDM3 |  | 0.9158% | -14.04% | -36.82% | 3.788% | 4.225% | 4.552% | 0.4633% | 6.023% | 3.298% | 5.128% | 6.643% | -20.52% | 1.281% | 10.11% | 12.42% | -15.19% | -14.53% | -17% | 3.134% | -0.8449% | -11.8% | -30.77% | 9.032%
| Z-Score (Absolute) |  | 0.05 | 1.93 | 4.79 | 0.32 | 0.37 | 0.41 | 0.1 | 0.6 | 0.25 | 0.48 | 0.67 | 2.74 | 0 | 1.11 | 1.4 | 2.07 | 1.99 | 2.3 | 0.23 | 0.27 | 1.65 | 4.03 | 0.97
| Analytical results for Ta in OREAS 610 (Certified Value 0.68 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.64 | 0.6 | 0.2 | 0.66 | 0.59 | 0.63 | 0.65 | 0.64 | 0.68 | 0.71 | 0.5165 | 0.8 | 0.76 | 0.7 | < 10 | 0.66 | 0.75 | 0.76 | 0.8 |  |  |  |  | 1
| 1 | 2 | 0.67 | 0.6 | 0.3 | 0.67 | 0.65 | 0.7 | 0.63 | 0.66 | 0.71 | 0.7 | 0.4747 | 0.7 | 0.73 | 0.8 | < 10 | 0.66 | 0.71 | 0.74 | 0.7 |  |  |  |  | 9
| 1 | 3 | 0.66 | 0.6 | 0.3 | 0.7 | 0.63 | 0.66 | 0.64 | 0.63 | 0.66 | 0.66 | 0.42884 | 0.7 | 0.69 | 0.8 | < 10 | 0.66 | 0.71 | 0.72 | 0.8 |  |  |  |  | 16
| 1 | 4 | 0.65 | 0.6 | 0.2 | 0.67 | 0.62 | 0.68 | 0.64 | 0.62 | 0.68 | 0.68 | 0.5177 | 0.8 | 0.76 | 0.8 | < 10 | 0.66 | 0.72 | 0.75 | 0.7 |  |  |  |  | 0.67739
| 1 | 5 | 0.64 | 0.6 | < 0.1 | 0.73 | 0.61 | 0.69 | 0.62 | 0.62 | 0.71 | 0.7 | 0.5451 | 0.8 | 0.75 | 0.8 | < 10 | 0.66 | 0.72 | 0.71 | 0.7 |  |  |  |  | 55
| 1 | 6 | 0.65 | 0.6 | < 0.1 | 0.66 | 0.63 | 0.67 | 0.62 | 0.67 | 0.71 | 0.67 | 0.47349 | 0.8 | 0.73 | 0.8 | < 10 | 0.66 | 0.73 | 0.76 | 0.7
| Mean |  | 0.65167 | 0.6 | 0.25 | 0.68167 | 0.62167 | 0.67167 | 0.63333 | 0.64 | 0.69167 | 0.68667 | 0.49272 | 0.76667 | 0.73667 | 0.78333 |  | 0.66 | 0.72333 | 0.74 | 0.73333
| Median |  | 0.65 | 0.6 | 0.25 | 0.67 | 0.625 | 0.675 | 0.635 | 0.635 | 0.695 | 0.69 | 0.4956 | 0.8 | 0.74 | 0.8 |  | 0.66 | 0.72 | 0.745 | 0.7
| Std Dev. |  | 0.01169 | 0 | 0.057735 | 0.027869 | 0.020412 | 0.024833 | 0.012111 | 0.020976 | 0.02137 | 0.019664 | 0.041735 | 0.05164 | 0.026583 | 0.040825 |  | 0 | 0.015055 | 0.020976 | 0.05164
| Rel.Std.Dev. |  | 1.794% | 0% | 23.09% | 4.088% | 3.283% | 3.697% | 1.912% | 3.278% | 3.09% | 2.864% | 8.47% | 6.736% | 3.609% | 5.212% |  | 0% | 2.081% | 2.835% | 7.042%
| PDM3 |  | -3.797% | -11.42% | -63.09% | 0.6315% | -8.226% | -0.8447% | -6.504% | -5.52% | 2.108% | 1.37% | -27.26% | 13.18% | 8.751% | 15.64% |  | -2.567% | 6.783% | 9.243% | 8.259%
| Z-Score (Absolute) |  | 0.42 | NA | NA | 0.08 | 0.93 | 0.08 | 0.73 | 0.62 | 0.25 | 0.17 | 3.1 | NA | 1.01 | NA | 72.88 | 0.28 | 0.79 | 1.07 | NA
Results from laboratories 3, 4, 14, 16 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 610 (Certified Value 0.31 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16; 2.07: 2.18; 2.08: 2.20; 2.09: 2.21 |  |  |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.38; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.26; 2.09: 0.33 |  |  |  |  | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.25; 2.09: 0.35 |  |  |  |  | 10
| 1 | 3 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.36; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.26; 2.09: 0.35 |  |  |  |  | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.36; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.26; 2.09: 0.34 |  |  |  |  | 0.30958
| 1 | 5 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.4; 2.06: 0.4; 2.07: 0.3; 2.08: 0.25; 2.09: 0.32 |  |  |  |  | 56
| 1 | 6 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.24; 2.09: 0.34
| Mean |  | 2.02: 0.2; 2.03: 0.3; 2.04: 0.36833; 2.05: 0.31667; 2.06: 0.4; 2.07: 0.3; 2.08: 0.25333; 2.09: 0.33833
| Median |  | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.255; 2.09: 0.34
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 0; 2.04: 0.0075277; 2.05: 0.040825; 2.06: 6.0809e-17; 2.07: 0; 2.08: 0.008165; 2.09: 0.01169
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 0%; 2.04: 2.044%; 2.05: 12.89%; 2.06: 1.52e-14%; 2.07: 0%; 2.08: 3.223%; 2.09: 3.455%
| PDM3 |  | 2.02: -35.4%; 2.03: -3.096%; 2.04: 18.98%; 2.05: 2.288%; 2.06: 29.21%; 2.07: -3.096%; 2.08: -18.17%; 2.09: 9.287%
| Z-Score (Absolute) |  | 2.02: 1.72; 2.03: 0.13; 2.04: 0.95; 2.05: 0.13; 2.06: 1.45; 2.07: 0.13; 2.08: 0.87; 2.09: 0.48
| Analytical results for Te in OREAS 610 (Certified Value 41.6 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 37.8 | 41.45 | 28.3 | 40.5 | 41.7 | 39.3 | 43.1 | 46.2 | 39.3 | 43.3 | 40.1367 | 41.41 | 42.6 | 29.6 | 49.7931 | 36.84 | 42.8 | 39.83 | 41.6 |  |  |  |  | 1
| 1 | 2 | 37.5 | 41.97 | 28.6 | 41.4 | 42.5 | 41.8 | 43 | 47.5 | 40.4 | 42.3 | 41.5643 | 39.02 | 41.2 | 30.4 | 48.6981 | 37.91 | 44 | 40.73 | 42.3 |  |  |  |  | 32
| 1 | 3 | 37.4 | 42.98 | 31 | 43.2 | 43 | 40.6 | 42.2 | 45.1 | 39.2 | 43.1 | 42.9778 | 41 | 42.1 | 30.6 | 45.7358 | 37.47 | 44.1 | 40.9 | 40.1 |  |  |  |  | 16
| 1 | 4 | 37.4 | 43.98 | 28.9 | 41.7 | 43.1 | 41.7 | 43.7 | 45.9 | 40 | 42.4 | 40.4047 | 43.51 | 43.6 | 30 | 49.6484 | 39.25 | 42.6 | 41.77 | 41.5 |  |  |  |  | 41.5506
| 1 | 5 | 37.9 | 43.13 | 24.4 | 42.9 | 41.3 | 41.6 | 41.7 | 44.9 | 38.2 | 42.8 | 42.7645 | 42.52 | 41.9 | 30.8 | 46.6207 | 37.94 | 42.3 | 39.79 | 41.3 |  |  |  |  | 57
| 1 | 6 | 38.5 | 42.42 | 21.5 | 41.5 | 42.5 | 40.6 | 41.9 | 47.3 | 38.2 | 44.3 | 42.5702 | 43.87 | 41.2 | 30 | 46.4025 | 37.77 | 43.1 | 39.08 | 40.2
| Mean |  | 37.75 | 42.655 | 27.1167 | 41.8667 | 42.35 | 40.9333 | 42.6 | 46.15 | 39.2167 | 43.0333 | 41.7364 | 41.8883 | 42.1 | 30.2333 | 47.8164 | 37.8633 | 43.15 | 40.35 | 41.1667
| Median |  | 37.65 | 42.7 | 28.45 | 41.6 | 42.5 | 41.1 | 42.6 | 46.05 | 39.25 | 42.95 | 42.0673 | 41.965 | 42 | 30.2 | 47.6594 | 37.84 | 42.95 | 40.28 | 41.4
| Std Dev. |  | 0.42308 | 0.90148 | 3.4868 | 1.0093 | 0.71484 | 0.96678 | 0.78486 | 1.084 | 0.90425 | 0.73121 | 1.2374 | 1.8001 | 0.91214 | 0.44572 | 1.7776 | 0.79266 | 0.74498 | 0.96542 | 0.85713
| Rel.Std.Dev. |  | 1.121% | 2.113% | 12.86% | 2.411% | 1.688% | 2.362% | 1.842% | 2.349% | 2.306% | 1.699% | 2.965% | 4.297% | 2.167% | 1.474% | 3.717% | 2.093% | 1.726% | 2.393% | 2.082%
| PDM3 |  | -9.147% | 2.658% | -34.74% | 0.7607% | 1.924% | -1.486% | 2.526% | 11.07% | -5.617% | 3.568% | 0.4471% | 0.8128% | 1.322% | -27.24% | 15.08% | -8.874% | 3.849% | -2.89% | -0.924%
| Z-Score (Absolute) |  | 2.38 | 0.46 | 8.53 | 0 | 0.28 | 0.54 | 0.42 | 2.48 | 1.53 | 0.67 | 0.08 | 0.01 | 0.13 | 6.72 | 3.44 | 2.31 | 0.74 | 0.88 | 0.4
| Analytical results for Th in OREAS 610 (Certified Value 8.95 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 9.07 | 5.8 | 2.3 | 8.65 | 8.58 | 9.18 | 8.65 | 10.3 | 7.98 | 9.19 | 8.2546 | 5.3 | 9.2 | 9.8 | < 10 | 7.48 | 5.4 | 8.5 | 10.6 |  |  |  |  | 1
| 1 | 2 | 9.08 | 6.6 | 2.3 | 9 | 8.68 | 8.9 | 8.54 | 10.15 | 7.82 | 9.73 | 8.4181 | 5.7 | 9.1 | 9.6 | < 10 | 8.58 | 5.4 | 8.4 | 10.3 |  |  |  |  | 33
| 1 | 3 | 9.19 | 6.5 | 2.3 | 9.16 | 8.86 | 8.79 | 8.91 | 9.48 | 8.03 | 9.09 | 8.5253 | 5.1 | 8.9 | 9.8 | 10.5804 | 8.41 | 5.4 | 7.6 | 10.3 |  |  |  |  | 16
| 1 | 4 | 9.61 | 6.3 | 2.2 | 8.94 | 8.81 | 9.2 | 8.68 | 9.93 | 8.53 | 9.46 | 8.5893 | 5.3 | 8.1 | 9.9 | 11.767 | 8.39 | 5.5 | 8.6 | 10.4 |  |  |  |  | 8.9466
| 1 | 5 | 9.29 | 6.3 | 3.9 | 8.79 | 8.45 | 8.92 | 8.49 | 10.3 | 7.95 | 9.01 | 8.3554 | 7.1 | 8.5 | 9.9 | < 10 | 8.35 | 5.3 | 7.9 | 10.1 |  |  |  |  | 58
| 1 | 6 | 9.21 | 6.3 | 3.7 | 9.34 | 8.87 | 8.8 | 8.11 | 11.25 | 8.1 | 9.59 | 8.1165 | 6.7 | 8.1 | 10 | 10.671 | 8.3 | 5.5 | 8.9 | 10.2
| Mean |  | 9.2417 | 6.3 | 2.7833 | 8.98 | 8.7083 | 8.965 | 8.5633 | 10.235 | 8.0683 | 9.345 | 8.3765 | 5.8667 | 8.65 | 9.8333 | 11.0061 | 8.2517 | 5.4167 | 8.3167 | 10.3167
| Median |  | 9.2 | 6.3 | 2.3 | 8.97 | 8.745 | 8.91 | 8.595 | 10.225 | 8.005 | 9.325 | 8.3868 | 5.5 | 8.7 | 9.85 | 10.671 | 8.37 | 5.4 | 8.45 | 10.3
| Std Dev. |  | 0.19864 | 0.27568 | 0.79099 | 0.2486 | 0.1694 | 0.18196 | 0.26561 | 0.58463 | 0.24457 | 0.29077 | 0.1744 | 0.83347 | 0.48888 | 0.13663 | 0.66049 | 0.38974 | 0.075277 | 0.47924 | 0.17224
| Rel.Std.Dev. |  | 2.149% | 4.376% | 28.42% | 2.768% | 1.945% | 2.03% | 3.102% | 5.712% | 3.031% | 3.112% | 2.082% | 14.21% | 5.652% | 1.389% | 6.001% | 4.723% | 1.39% | 5.762% | 1.67%
| PDM3 |  | 3.298% | -29.58% | -68.89% | 0.373% | -2.663% | 0.2054% | -4.284% | 14.4% | -9.817% | 4.453% | -6.372% | -34.43% | -3.315% | 9.911% | 23.02% | -7.768% | -39.46% | -7.041% | 15.31%
| Z-Score (Absolute) |  | 0.82 | 2.72 | 6.96 | 0.5 | 0.17 | 0.48 | 0 | 2.01 | 0.6 | 0.94 | 0.22 | 3.25 | 0.1 | 1.53 | 0.67 | 0.38 | 3.79 | 0.3 | 2.11
| Analytical results for Ti in OREAS 610 (Certified Value 0.167 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.1638 | 0.167 | 0.17 | 0.166 | 0.166 | 0.155 | 0.163 | 0.168 | 0.164 | 0.1746 | 0.17518 | 0.1667 | 0.17 | 0.18 | 0.16017 | 0.16 | 0.156 | 0.16 | 0.17 | 0.15 | 0.167 | 0.17 | 0.1677 | 1
| 1 | 2 | 0.1609 | 0.17 | 0.171 | 0.17 | 0.164 | 0.166 | 0.163 | 0.173 | 0.163 | 0.1701 | 0.17672 | 0.1667 | 0.17 | 0.18 | 0.16008 | 0.15 | 0.159 | 0.15 | 0.17 | 0.15 | 0.167 | 0.17 | 0.1706 | 11
| 1 | 3 | 0.1629 | 0.167 | 0.168 | 0.168 | 0.168 | 0.163 | 0.163 | 0.165 | 0.161 | 0.1699 | 0.17847 | 0.1676 | 0.16 | 0.19 | 0.16267 | 0.15 | 0.161 | 0.14 | 0.16 | 0.15 | 0.165 | 0.16 | 0.1735 | 16
| 1 | 4 | 0.1615 | 0.17 | 0.172 | 0.168 | 0.168 | 0.166 | 0.165 | 0.167 | 0.164 | 0.1753 | 0.17954 | 0.1676 | 0.17 | 0.18 | 0.15939 | 0.16 | 0.159 | 0.15 | 0.17 | 0.16 | 0.168 | 0.16 | 0.1744 | 0.16664
| 1 | 5 | 0.1644 | 0.164 | 0.175 | 0.167 | 0.167 | 0.166 | 0.161 | 0.168 | 0.168 | 0.175 | 0.17591 | 0.1751 | 0.18 | 0.18 | 0.16378 | 0.16 | 0.162 | 0.14 | 0.17 | 0.15 | 0.171 | 0.17 | 0.1693 | 59
| 1 | 6 | 0.1615 | 0.173 | 0.17 | 0.165 | 0.166 | 0.165 | 0.162 | 0.173 | 0.162 | 0.1719 | 0.17923 | 0.1722 | 0.17 | 0.19 | 0.16243 | 0.16 | 0.156 | 0.15 | 0.17 | 0.15 | 0.169 | 0.16 | 0.1622
| Mean |  | 0.1625 | 0.1685 | 0.171 | 0.16733 | 0.1665 | 0.1635 | 0.16283 | 0.169 | 0.16367 | 0.1728 | 0.17751 | 0.16932 | 0.17 | 0.18333 | 0.16142 | 0.15667 | 0.15883 | 0.14833 | 0.16833 | 0.15167 | 0.16783 | 0.165 | 0.16962
| Median |  | 0.1622 | 0.1685 | 0.1705 | 0.1675 | 0.1665 | 0.1655 | 0.163 | 0.168 | 0.1635 | 0.17325 | 0.1776 | 0.1676 | 0.17 | 0.18 | 0.1613 | 0.16 | 0.159 | 0.15 | 0.17 | 0.15 | 0.1675 | 0.165 | 0.16995
| Std Dev. |  | 0.0014156 | 0.0031464 | 0.0023664 | 0.0017512 | 0.0015166 | 0.0043243 | 0.0013292 | 0.0032863 | 0.0024221 | 0.0024835 | 0.0018206 | 0.0035028 | 0.0063246 | 0.005164 | 0.0017681 | 0.005164 | 0.0024833 | 0.0075277 | 0.0040825 | 0.0040825 | 0.0020412 | 0.0054772 | 0.0044206
| Rel.Std.Dev. |  | 0.8712% | 1.867% | 1.384% | 1.047% | 0.9109% | 2.645% | 0.8163% | 1.945% | 1.48% | 1.437% | 1.026% | 2.069% | 3.72% | 2.817% | 1.095% | 3.296% | 1.563% | 5.075% | 2.425% | 2.692% | 1.216% | 3.32% | 2.606%
| PDM3 |  | -2.482% | 1.119% | 2.619% | 0.4189% | -0.08121% | -1.882% | -2.282% | 1.419% | -1.782% | 3.7% | 6.525% | 1.609% | 2.019% | 10.02% | -3.129% | -5.982% | -4.682% | -10.98% | 1.019% | -8.983% | 0.7189% | -0.9814% | 1.789%
| Z-Score (Absolute) |  | 0.89 | 0.21 | 0.67 | 0 | 0.15 | 0.7 | 0.83 | 0.31 | 0.67 | 1.01 | 1.87 | 0.36 | 0.49 | 2.94 | 1.09 | 1.96 | 1.56 | 3.49 | 0.18 | 2.88 | 0.09 | 0.43 | 0.42
| Analytical results for Tl in OREAS 610 (Certified Value 1.84 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.75 | 1.79 | 1.96 | 1.66 | 1.9 | 1.63 | 1.82 | 1.82 | 1.78 | 1.94 | 2.0527 | 1.9 | 1.65 | 1.8 | < 5 | 1.5 | 1.73 | 1.64 | 2.2 |  |  |  |  | 1
| 1 | 2 | 1.83 | 1.83 | 1.98 | 1.79 | 1.88 | 1.83 | 1.82 | 1.82 | 1.85 | 1.89 | 2.0002 | 1.92 | 1.61 | 1.9 | < 5 | 1.5 | 1.78 | 1.68 | 2.2 |  |  |  |  | 35
| 1 | 3 | 1.79 | 1.79 | 1.97 | 1.81 | 1.87 | 1.83 | 1.84 | 1.73 | 1.82 | 1.83 | 1.9984 | 2.01 | 1.69 | 1.9 | < 5 | 1.5 | 1.77 | 1.72 | 2.2 |  |  |  |  | 16
| 1 | 4 | 1.78 | 1.92 | 1.95 | 1.71 | 1.91 | 1.86 | 1.87 | 1.78 | 1.8 | 1.89 | 2.0016 | 2.02 | 1.7 | 1.8 | < 5 | 1.5 | 1.77 | 1.65 | 2.1 |  |  |  |  | 1.8371
| 1 | 5 | 1.82 | 1.85 | 2.01 | 1.85 | 1.81 | 1.82 | 1.78 | 1.76 | 1.81 | 1.92 | 1.9976 | 2.16 | 1.66 | 1.9 | < 5 | 1.5 | 1.81 | 1.65 | 2.1 |  |  |  |  | 60
| 1 | 6 | 1.78 | 1.9 | 2.04 | 1.78 | 1.89 | 1.81 | 1.8 | 1.88 | 1.75 | 1.88 | 2.0461 | 2.06 | 1.61 | 1.9 | < 5 | 1.4 | 1.76 | 1.64 | 2.1
| Mean |  | 1.7917 | 1.8467 | 1.985 | 1.7667 | 1.8767 | 1.7967 | 1.8217 | 1.7983 | 1.8017 | 1.8917 | 2.0161 | 2.0117 | 1.6533 | 1.8667 |  | 1.4833 | 1.77 | 1.6633 | 2.15
| Median |  | 1.785 | 1.84 | 1.975 | 1.785 | 1.885 | 1.825 | 1.82 | 1.8 | 1.805 | 1.89 | 2.0009 | 2.015 | 1.655 | 1.9 |  | 1.5 | 1.77 | 1.65 | 2.15
| Std Dev. |  | 0.029269 | 0.05465 | 0.033912 | 0.069474 | 0.03559 | 0.083347 | 0.031252 | 0.053072 | 0.034303 | 0.037639 | 0.025922 | 0.095167 | 0.038297 | 0.05164 |  | 0.040825 | 0.026077 | 0.031411 | 0.054772
| Rel.Std.Dev. |  | 1.634% | 2.959% | 1.708% | 3.933% | 1.896% | 4.639% | 1.716% | 2.951% | 1.904% | 1.99% | 1.286% | 4.731% | 2.316% | 2.766% |  | 2.752% | 1.473% | 1.888% | 2.548%
| PDM3 |  | -2.471% | 0.5225% | 8.053% | -3.832% | 2.156% | -2.199% | -0.8384% | -2.109% | -1.927% | 2.972% | 9.745% | 9.504% | -10% | 1.611% |  | -19.26% | -3.651% | -9.457% | 17.03%
| Z-Score (Absolute) |  | 0.37 | 0.31 | 2 | 0.67 | 0.67 | 0.31 | 0 | 0.29 | 0.25 | 0.86 | 2.38 | 2.33 | 2.06 | 0.55 | 8.32 | 4.15 | 0.63 | 1.94 | 4.03
| Analytical results for Tm in OREAS 610 (Indicative Value 91.7 ppb) | [24 empty] | Indicative
Tm | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 |  |  |  |  | ppb
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 80 |  |  |  |  | 1
| 1 | 2 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100 |  |  |  |  | 6
| 1 | 3 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100 |  |  |  |  | 16
| 1 | 4 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100 |  |  |  |  | 91.6667
| 1 | 5 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100 |  |  |  |  | 12
| 1 | 6 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100
| Mean |  | 2.04: 86.6667; 2.06: 96.6667
| Median |  | 2.04: 90; 2.06: 100
| Std Dev. |  | 2.04: 5.164; 2.06: 8.165
| Rel.Std.Dev. |  | 2.04: 5.958%; 2.06: 8.447%
| PDM3 |  | 2.04: -5.455%; 2.06: 5.455%
Z-Score (Absolute) | bar: 2.02 to 2.06
| Analytical results for U in OREAS 610 (Certified Value 2.63 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.57 | 2.6 | 2.8 | 2.6 | 2.6 | 2.6 | 2.5 | 2.5 | 2.67 | 2.86 | 2.6914 | 2.9 | 2.65 | 2.6 | < 10 | 1.93 | 2.5 | 2.44 | 3.4 |  |  |  |  | 1
| 1 | 2 | 2.56 | 2.7 | 2.8 | 2.7 | 2.6 | 2.6 | 2.5 | 2.5 | 2.69 | 2.78 | 2.7475 | 2.8 | 2.66 | 2.6 | < 10 | 1.97 | 2.6 | 2.5 | 3.3 |  |  |  |  | 37
| 1 | 3 | 2.67 | 2.6 | 2.7 | 2.9 | 2.5 | 2.6 | 2.6 | 2.5 | 2.57 | 2.77 | 2.7887 | 2.6 | 2.62 | 2.6 | < 10 | 1.94 | 2.7 | 2.5 | 3.1 |  |  |  |  | 16
| 1 | 4 | 2.56 | 2.6 | 2.8 | 2.6 | 2.6 | 2.7 | 2.6 | 2.5 | 2.61 | 2.76 | 2.6839 | 2.8 | 2.65 | 2.6 | < 10 | 2 | 2.5 | 2.45 | 3.1 |  |  |  |  | 2.6326
| 1 | 5 | 2.61 | 2.8 | 2.9 | 2.7 | 2.5 | 2.6 | 2.5 | 2.5 | 2.5 | 2.78 | 2.7495 | 3.1 | 2.65 | 2.7 | < 10 | 1.99 | 2.5 | 2.4 | 3.1 |  |  |  |  | 61
| 1 | 6 | 2.63 | 2.7 | 2.8 | 2.7 | 2.6 | 2.6 | 2.4 | 2.6 | 2.51 | 2.79 | 2.6782 | 3.2 | 2.59 | 2.6 | < 10 | 1.92 | 2.5 | 2.44 | 3.1
| Mean |  | 2.6 | 2.6667 | 2.8 | 2.7 | 2.5667 | 2.6167 | 2.5167 | 2.5167 | 2.5917 | 2.79 | 2.7232 | 2.9 | 2.6367 | 2.6167 |  | 1.9583 | 2.55 | 2.455 | 3.1833
| Median |  | 2.59 | 2.65 | 2.8 | 2.7 | 2.6 | 2.6 | 2.5 | 2.5 | 2.59 | 2.78 | 2.7194 | 2.85 | 2.65 | 2.6 |  | 1.955 | 2.5 | 2.445 | 3.1
| Std Dev. |  | 0.044721 | 0.08165 | 0.063246 | 0.10954 | 0.05164 | 0.040825 | 0.075277 | 0.040825 | 0.079603 | 0.035777 | 0.045065 | 0.21909 | 0.026583 | 0.040825 |  | 0.033116 | 0.083666 | 0.038859 | 0.13292
| Rel.Std.Dev. |  | 1.72% | 3.062% | 2.259% | 4.057% | 2.012% | 1.56% | 2.991% | 1.622% | 3.072% | 1.282% | 1.655% | 7.555% | 1.008% | 1.56% |  | 1.691% | 3.281% | 1.583% | 4.175%
| PDM3 |  | -1.238% | 1.294% | 6.359% | 2.56% | -2.504% | -0.605% | -4.404% | -4.404% | -1.555% | 5.979% | 3.441% | 10.16% | 0.1547% | -0.605% |  | -25.61% | -3.137% | -6.746% | 20.92%
| Z-Score (Absolute) |  | 0.11 | 0.34 | 1.24 | 0.56 | 0.34 | 0 | 0.67 | 0.67 | 0.17 | 1.17 | 0.72 | 1.91 | 0.13 | 0 | 16.07 | 4.44 | 0.45 | 1.09 | 3.82
| Analytical results for V in OREAS 610 (Certified Value 30.5 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 31 | 30 | 28 | 30 | 30 | 29 | 30 | 31 | 30 | 33 | 32.0042 | 27 | 31.6 | 35 | 30.7492 | 29 | 31.238 | 31 | 31 | 27 | 34 | 31 | 35.8 | 1
| 1 | 2 | 29 | 31 | 29 | 31 | 30 | 30 | 30 | 32 | 31 | 33 | 31.3776 | 27 | 31.5 | 40 | 30.6868 | 27 | 30.679 | 32 | 32 | 26 | 34 | 31 | 35.94 | 38
| 1 | 3 | 31 | 31 | 28 | 30 | 30 | 30 | 30 | 30 | 31 | 33 | 30.5024 | 28 | 33.3 | 35 | 30.7636 | 28 | 32.936 | 32 | 31 | 27 | 34 | 31 | 36.08 | 16
| 1 | 4 | 31 | 30 | 29 | 30 | 30 | 30 | 30 | 31 | 31 | 32 | 30.087 | 26 | 32.3 | 35 | 30.6942 | 29 | 30.658 | 30 | 30 | 27 | 34 | 31 | 36.22 | 30.4801
| 1 | 5 | 31 | 31 | 30 | 30 | 30 | 31 | 29 | 31 | 30 | 32 | 29.9592 | 33 | 32 | 40 | 30.8189 | 28 | 31.44 | 31 | 30 | 26 | 34 | 31 | 36.15 | 62
| 1 | 6 | 31 | 31 | 28 | 29 | 29 | 30 | 30 | 32 | 30 | 33 | 31.8514 | 33 | 31.7 | 35 | 30.8194 | 27 | 30.898 | 31 | 31 | 26 | 34 | 31 | 36.15
| Mean |  | 30.6667 | 30.6667 | 28.6667 | 30 | 29.8333 | 30 | 29.8333 | 31.1667 | 30.5 | 32.6667 | 30.9636 | 29 | 32.0667 | 36.6667 | 30.7554 | 28 | 31.3082 | 31.1667 | 30.8333 | 26.5 | 34 | 31 | 36.0567
| Median |  | 31 | 31 | 28.5 | 30 | 30 | 30 | 30 | 31 | 30.5 | 33 | 30.94 | 27.5 | 31.85 | 35 | 30.7564 | 28 | 31.068 | 31 | 31 | 26.5 | 34 | 31 | 36.115
| Std Dev. |  | 0.8165 | 0.5164 | 0.8165 | 0.63246 | 0.40825 | 0.63246 | 0.40825 | 0.75277 | 0.54772 | 0.5164 | 0.89803 | 3.1623 | 0.67132 | 2.582 | 0.057763 | 0.89443 | 0.85545 | 0.75277 | 0.75277 | 0.54772 | 0 | 0 | 0.15756
| Rel.Std.Dev. |  | 2.662% | 1.684% | 2.848% | 2.108% | 1.368% | 2.108% | 1.368% | 2.415% | 1.796% | 1.581% | 2.9% | 10.9% | 2.094% | 7.042% | 0.1878% | 3.194% | 2.732% | 2.415% | 2.441% | 2.067% | 0% | 0% | 0.437%
| PDM3 |  | 0.6122% | 0.6122% | -5.95% | -1.575% | -2.122% | -1.575% | -2.122% | 2.253% | 0.06534% | 7.174% | 1.586% | -4.856% | 5.205% | 20.3% | 0.9031% | -8.137% | 2.717% | 2.253% | 1.159% | -13.06% | 11.55% | 1.706% | 18.3%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 1.86 | 0.67 | 0.82 | 0.67 | 0.82 | 0.37 | 0.23 | 1.71 | 0.19 | 1.57 | 1.17 | 5.28 | 0 | 2.46 | 0.49 | 0.37 | 0.07 | 3.8 | 2.9 | 0.22 | 4.73
| Analytical results for W in OREAS 610 (Certified Value 7.57 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 7.5 | 7 | 6.5 | 7.4 | 7.5 | 7 | 7.6 | 7.6 | 7.2 | 8 | 7.1229 | 8.1 | 8.5 | 7.5 | < 20 | 7.1 | < 10 | 7.6 | 7.3 | 13 | 8.9 | < 5 |  | 1
| 1 | 2 | 7.7 | 7 | 6.5 | 7.7 | 7.8 | 7.4 | 7.3 | 8.1 | 7.5 | 7.9 | 7.212 | 7.8 | 8.8 | 7.5 | < 20 | 6.8 | < 10 | 8 | 7.3 | 12 | 8.9 | < 5 |  | 39
| 1 | 3 | 7.7 | 7 | 6 | 8 | 7.9 | 7.3 | 7.3 | 7.5 | 7.1 | 7.8 | 7.257 | 8.2 | 8.3 | 7.5 | < 20 | 6.9 | < 10 | 7.6 | 7.3 | 13 | 9 | < 5 |  | 16
| 1 | 4 | 7.6 | 7.5 | 5.8 | 7.5 | 7.7 | 7.6 | 7.7 | 7.7 | 7.2 | 7.6 | 7.1705 | 8.3 | 8.7 | 7.5 | < 20 | 7.1 | < 10 | 8 | 7.2 | 13 | 8.8 | < 5 |  | 7.5661
| 1 | 5 | 7.6 | 7 | 6 | 7.7 | 7.7 | 7.6 | 7.1 | 7.4 | 7.6 | 7.8 | 7.3267 | 8.1 | 8.3 | 7.5 | < 20 | 7.1 | < 10 | 7.9 | 7.1 | 12 | 8.6 | < 5 |  | 63
| 1 | 6 | 7.4 | 8.2 | 5.5 | 7.5 | 7.5 | 7.4 | 7.5 | 8.1 | 7.4 | 7.8 | 7.2574 | 9.1 | 8.6 | 7.5 | < 20 | 6.6 | < 10 | 8 | 7.2 | 13 | 8.6 | < 5
| Mean |  | 7.5833 | 7.2833 | 6.05 | 7.6333 | 7.6833 | 7.3833 | 7.4167 | 7.7333 | 7.3333 | 7.8167 | 7.2244 | 8.2667 | 8.5333 | 7.5 |  | 6.9333 |  | 7.85 | 7.2333 | 12.6667 | 8.8
| Median |  | 7.6 | 7 | 6 | 7.6 | 7.7 | 7.4 | 7.4 | 7.65 | 7.3 | 7.8 | 7.2345 | 8.15 | 8.55 | 7.5 |  | 7 |  | 7.95 | 7.25 | 13 | 8.85
| Std Dev. |  | 0.1169 | 0.4916 | 0.3937 | 0.21602 | 0.16021 | 0.22286 | 0.22286 | 0.30111 | 0.19664 | 0.13292 | 0.072084 | 0.44121 | 0.20656 | 0 |  | 0.20656 |  | 0.19748 | 0.08165 | 0.5164 | 0.16733
| Rel.Std.Dev. |  | 1.542% | 6.75% | 6.507% | 2.83% | 2.085% | 3.018% | 3.005% | 3.894% | 2.681% | 1.7% | 0.9978% | 5.337% | 2.421% | 0% |  | 2.979% |  | 2.516% | 1.129% | 4.077% | 1.902%
| PDM3 |  | 0.2276% | -3.737% | -20.04% | 0.8885% | 1.549% | -2.416% | -1.975% | 2.21% | -3.077% | 3.312% | -4.516% | 9.259% | 12.78% | -0.8738% |  | -8.363% |  | 3.752% | -4.398% | 67.41% | 16.31%
| Z-Score (Absolute) |  | 0.2 | 0.53 | 3.55 | 0.33 | 0.45 | 0.29 | 0.2 | 0.57 | 0.41 | 0.77 | 0.67 | 1.88 | 2.53 | 0 | 6.12 | 1.39 | 6.12 | 0.86 | 0.65 | NA | 3.18 | 12.23
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 610 (Certified Value 6.59 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 6.8 | 6.6 | 6.8 | 6.3 | 6.3 | 6.3 | 6.8 | 7.2 | 6.5 | 6.88 | 6.522 | 5.2 | 6.7 | 6.8 | 6.9088 | 8.8 | 5.868 | 5.1 | 6.4 | 5.9 | 7.1 | 5 |  | 1
| 1 | 2 | 6.8 | 6.7 | 6.7 | 6.4 | 6.3 | 6.9 | 6.7 | 7.5 | 6.3 | 6.81 | 6.363 | 5.5 | 6.6 | 6.8 | 6.9395 | 8.9 | 6.148 | 5.3 | 6.3 | 5.7 | 6.7 | 5 |  | 40
| 1 | 3 | 6.7 | 6.8 | 6.4 | 6.8 | 6.7 | 6.7 | 6.6 | 7 | 6.4 | 6.72 | 6.4636 | 5.5 | 6.6 | 7 | 7.0281 | 8.5 | 5.956 | 5.2 | 6.6 | 5.9 | 6.9 | 5 |  | 16
| 1 | 4 | 6.8 | 6.8 | 6.5 | 6.5 | 6.4 | 6.7 | 6.7 | 7.2 | 6.6 | 6.6 | 6.3507 | 5.4 | 6.9 | 7 | 7.0246 | 8.8 | 5.294 | 5 | 6.4 | 5.8 | 6.7 | 5 |  | 6.5911
| 1 | 5 | 6.8 | 6.4 | 6.8 | 6.6 | 6.2 | 6.8 | 6.5 | 7.1 | 6.5 | 6.72 | 6.2953 | 6.2 | 6.8 | 6.8 | 7.09 | 8.9 | 5.702 | 5.1 | 6.2 | 5.7 | 6.9 | 5 |  | 64
| 1 | 6 | 6.8 | 6.5 | 6.5 | 6.5 | 6.3 | 6.7 | 6.3 | 7.3 | 6.5 | 6.95 | 6.2901 | 6.1 | 6.6 | 7 | 7.0287 | 8.9 | 5.822 | 5 | 6.4 | 5.8 | 7.1 | 5
| Mean |  | 6.7833 | 6.6333 | 6.6167 | 6.5167 | 6.3667 | 6.6833 | 6.6 | 7.2167 | 6.4667 | 6.78 | 6.3808 | 5.65 | 6.7 | 6.9 | 7.0033 | 8.8 | 5.7983 | 5.1167 | 6.3833 | 5.8 | 6.9 | 5
| Median |  | 6.8 | 6.65 | 6.6 | 6.5 | 6.3 | 6.7 | 6.65 | 7.2 | 6.5 | 6.765 | 6.3569 | 5.5 | 6.65 | 6.9 | 7.0263 | 8.85 | 5.845 | 5.1 | 6.4 | 5.8 | 6.9 | 5
| Std Dev. |  | 0.040825 | 0.1633 | 0.17224 | 0.17224 | 0.17512 | 0.20412 | 0.17889 | 0.17224 | 0.10328 | 0.12602 | 0.093341 | 0.40373 | 0.12649 | 0.10954 | 0.066686 | 0.15492 | 0.2885 | 0.1169 | 0.13292 | 0.089443 | 0.17889 | 0
| Rel.Std.Dev. |  | 0.6018% | 2.462% | 2.603% | 2.643% | 2.751% | 3.054% | 2.71% | 2.387% | 1.597% | 1.859% | 1.463% | 7.146% | 1.888% | 1.588% | 0.9522% | 1.76% | 4.976% | 2.285% | 2.082% | 1.542% | 2.593% | 0%
| PDM3 |  | 2.916% | 0.6407% | 0.3878% | -1.129% | -3.405% | 1.399% | 0.1349% | 9.491% | -1.888% | 2.866% | -3.191% | -14.28% | 1.652% | 4.687% | 6.253% | 33.51% | -12.03% | -22.37% | -3.152% | -12% | 4.687% | -24.14%
| Z-Score (Absolute) |  | 0.48 | 0.05 | 0 | 0.29 | 0.71 | 0.19 | 0.05 | 1.72 | 0.43 | 0.47 | 0.67 | 2.76 | 0.24 | 0.81 | 1.11 | 6.24 | 2.34 | 4.29 | 0.67 | 2.33 | 0.81 | NA
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 610 (Certified Value 0.52 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16; 2.07: 2.18; 2.08: 2.20; 2.09: 2.21 |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  | 2
2
| 1 | 1 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.6; 2.05: 0.5; 2.06: 0.7; 2.07: 0.5; 2.08: 0.4; 2.09: 0.5 |  | 1
| 1 | 2 | 2.02: 0.6; 2.03: 0.5; 2.04: 0.58; 2.05: 0.6; 2.06: 0.7; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5 |  | 13
| 1 | 3 | 2.02: 0.5; 2.03: 0.5; 2.04: 0.57; 2.05: 0.5; 2.06: 0.7; 2.07: 0.5; 2.08: 0.4; 2.09: 0.5 |  | 16
| 1 | 4 | 2.02: 0.5; 2.03: 0.5; 2.04: 0.54; 2.05: 0.5; 2.06: 0.75; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5 |  | 0.51952
| 1 | 5 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.56; 2.05: 0.5; 2.06: 0.75; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5 |  | 65
| 1 | 6 | 2.02: 0.6; 2.03: 0.5; 2.04: 0.57; 2.05: 0.5; 2.06: 0.7; 2.07: 0.5; 2.08: 0.4; 2.09: 0.5
| Mean |  | 2.02: 0.56667; 2.03: 0.53333; 2.04: 0.57; 2.05: 0.51667; 2.06: 0.71667; 2.07: 0.5; 2.08: 0.45; 2.09: 0.5
| Median |  | 2.02: 0.6; 2.03: 0.5; 2.04: 0.57; 2.05: 0.5; 2.06: 0.7; 2.07: 0.5; 2.08: 0.45; 2.09: 0.5
| Std Dev. |  | 2.02: 0.05164; 2.03: 0.05164; 2.04: 0.02; 2.05: 0.040825; 2.06: 0.02582; 2.07: 0; 2.08: 0.054772; 2.09: 0
| Rel.Std.Dev. |  | 2.02: 9.113%; 2.03: 9.682%; 2.04: 3.509%; 2.05: 7.902%; 2.06: 3.603%; 2.07: 0%; 2.08: 12.17%; 2.09: 0%
| PDM3 |  | 2.02: 9.074%; 2.03: 2.658%; 2.04: 9.716%; 2.05: -0.55%; 2.06: 37.95%; 2.07: -3.758%; 2.08: -13.38%; 2.09: -3.758%
| Z-Score (Absolute) |  | 2.02: 0.84; 2.03: 0.17; 2.04: 0.91; 2.05: 0.17; 2.06: 3.88; 2.07: 0.51; 2.08: 1.52; 2.09: 0.51
| Analytical results for Zn in OREAS 610 (Certified Value 1754 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 1805 | 1695.2 | 1660 | 1700 | 1800 | 1600 | 1770 | 1810 | 1960 | 1759 | 1785.4991 | 1827.5 | 1800 | 1700 | 1778.6721 | 1715 | 1787.66 | 1632 | 1909 | 1755 | 1842 | 1740 | 1712 | 1
| 1 | 2 | 1746 | 1666.5 | 1680 | 1740 | 1790 | 1730 | 1740 | 1860 | 1940 | 1681 | 1786.0898 | 1831.5 | 1840 | 1630 | 1759.4339 | 1687 | 1777.454 | 1629 | 1966 | 1683 | 1850 | 1730 | 1670 | 22
| 1 | 3 | 1761 | 1681.3 | 1640 | 1730 | 1790 | 1710 | 1730 | 1780 | 1920 | 1736 | 1769.0319 | 1845.9 | 1730 | 1680 | 1775.8182 | 1692 | 1851.057 | 1663 | 1923 | 1735 | 1844 | 1730 | 1678 | 16
| 1 | 4 | 1738 | 1685 | 1680 | 1740 | 1800 | 1740 | 1760 | 1800 | 1960 | 1727 | 1798.7462 | 1842.3 | 1760 | 1620 | 1756.2771 | 1723 | 1804.352 | 1685 | 1920 | 1757 | 1856 | 1710 | 1686 | 1753.6095
| 1 | 5 | 1777 | 1690.5 | 1730 | 1720 | 1780 | 1740 | 1720 | 1780 | 1960 | 1756 | 1792.1462 | 2003 | 1860 | 1680 | 1797.1237 | 1726 | 1796.597 | 1624 | 1928 | 1640 | 1861 | 1750 | 1684 | 66
| 1 | 6 | 1753 | 1727.1 | 1630 | 1680 | 1770 | 1720 | 1730 | 1870 | 1890 | 1721 | 1776.7942 | 1978 | 1770 | 1630 | 1783.1432 | 1724 | 1803.121 | 1630 | 1904 | 1686 | 1872 | 1720 | 1702
| Mean |  | 1763.3333 | 1690.9333 | 1670 | 1718.3333 | 1788.3333 | 1706.6667 | 1741.6667 | 1816.6667 | 1938.3333 | 1730 | 1784.7179 | 1888.0333 | 1793.3333 | 1656.6667 | 1775.078 | 1711.1667 | 1803.3735 | 1643.8333 | 1925 | 1709.3333 | 1854.1667 | 1730 | 1688.6667
| Median |  | 1757 | 1687.75 | 1670 | 1725 | 1790 | 1725 | 1735 | 1805 | 1950 | 1731.5 | 1785.7945 | 1844.1 | 1785 | 1655 | 1777.2452 | 1719 | 1799.859 | 1631 | 1921.5 | 1710.5 | 1853 | 1730 | 1685
| Std Dev. |  | 24.4022 | 20.2554 | 35.7771 | 24.0139 | 11.6905 | 53.5413 | 19.4079 | 39.3277 | 28.5774 | 28.4394 | 10.6142 | 80.048 | 49.6655 | 33.8625 | 15.2495 | 17.2675 | 25.4535 | 24.5228 | 21.9818 | 47.0347 | 11.2857 | 14.1421 | 15.5778
| Rel.Std.Dev. |  | 1.384% | 1.198% | 2.142% | 1.398% | 0.6537% | 3.137% | 1.114% | 2.165% | 1.474% | 1.644% | 0.5947% | 4.24% | 2.769% | 2.044% | 0.8591% | 1.009% | 1.411% | 1.492% | 1.142% | 2.752% | 0.6087% | 0.8175% | 0.9225%
| PDM3 |  | 0.5545% | -3.574% | -4.768% | -2.012% | 1.98% | -2.677% | -0.681% | 3.596% | 10.53% | -1.346% | 1.774% | 7.666% | 2.265% | -5.528% | 1.224% | -2.42% | 2.838% | -6.26% | 9.774% | -2.525% | 5.734% | -1.346% | -3.703%
| Z-Score (Absolute) |  | 0.29 | 0.67 | 0.95 | 0.31 | 0.62 | 0.47 | 0 | 1 | 2.61 | 0.16 | 0.57 | 1.95 | 0.69 | 1.13 | 0.44 | 0.41 | 0.82 | 1.3 | 2.44 | 0.43 | 1.5 | 0.16 | 0.7
| Analytical results for Zr in OREAS 610 (Certified Value 61 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 56.6 | 54 | 72 | 57 | 60.3 | 58.2 | 61.7 | 63.7 | 59.2 | 68.7 | 60.3593 | 61.9 | 63.2 | 61 | 65.8472 | 56 | 53.391 | 63.1 | 57.1 | 53.7 | 62.1 | 52 | 64 | 1
| 1 | 2 | 56 | 56 | 69 | 58.7 | 60.7 | 61 | 61.4 | 65.9 | 61.8 | 59.8 | 59.9236 | 62.7 | 62 | 63 | 70.204 | 55 | 54.593 | 64.1 | 56.1 | 53.3 | 61.1 | 51 | 68.65 | 23
| 1 | 3 | 55.2 | 56.7 | 70 | 63 | 63 | 59.5 | 60.9 | 62.4 | 61.9 | 61 | 59.784 | 63.4 | 60.5 | 63 | 70.0095 | 56 | 59.751 | 63.2 | 55.6 | 54.4 | 59.5 | 51 | 69.3 | 16
| 1 | 4 | 55.6 | 55.6 | 70 | 59 | 62 | 59.1 | 60.7 | 63.5 | 68.1 | 63 | 60.4393 | 63.8 | 65.2 | 64 | 69.0717 | 55 | 51.151 | 61.2 | 53.1 | 53.5 | 61.4 | 51 | 69.95 | 60.6013
| 1 | 5 | 56.6 | 55.8 | 75 | 59.3 | 59 | 60.8 | 60.2 | 61.7 | 67 | 61.5 | 60.3903 | 64 | 62.7 | 69 | 66.5927 | 55 | 57.532 | 61.4 | 54.4 | 52.8 | 62.6 | 51 | 70.6 | 67
| 1 | 6 | 55.4 | 56.2 | 68 | 58.2 | 62.5 | 59.7 | 58.8 | 65.2 | 69.5 | 62.8 | 60.9452 | 57 | 61 | 64 | 66.3851 | 54 | 53.783 | 65.6 | 55.3 | 54.7 | 61 | 52 | 70.25
| Mean |  | 55.9 | 55.7167 | 70.6667 | 59.2 | 61.25 | 59.7167 | 60.6167 | 63.7333 | 64.5833 | 62.8 | 60.3069 | 62.1333 | 62.4333 | 64 | 68.0184 | 55.1667 | 55.0335 | 63.1 | 55.2667 | 53.7333 | 61.2833 | 51.3333 | 68.7917
| Median |  | 55.8 | 55.9 | 70 | 58.85 | 61.35 | 59.6 | 60.8 | 63.6 | 64.45 | 62.15 | 60.3748 | 63.05 | 62.35 | 63.5 | 67.8322 | 55 | 54.188 | 63.15 | 55.45 | 53.6 | 61.25 | 51 | 69.625
| Std Dev. |  | 0.60332 | 0.92177 | 2.5033 | 2.0288 | 1.5136 | 1.0534 | 1.0342 | 1.6008 | 4.1547 | 3.1235 | 0.41353 | 2.6303 | 1.6908 | 2.6833 | 1.9629 | 0.75277 | 3.0992 | 1.6589 | 1.3866 | 0.70616 | 1.0685 | 0.5164 | 2.4486
| Rel.Std.Dev. |  | 1.079% | 1.654% | 3.542% | 3.427% | 2.471% | 1.764% | 1.706% | 2.512% | 6.433% | 4.974% | 0.6857% | 4.233% | 2.708% | 4.193% | 2.886% | 1.365% | 5.632% | 2.629% | 2.509% | 1.314% | 1.744% | 1.006% | 3.559%
| PDM3 |  | -7.758% | -8.06% | 16.61% | -2.312% | 1.071% | -1.46% | 0.02543% | 5.168% | 6.571% | 3.628% | -0.4856% | 2.528% | 3.023% | 5.608% | 12.24% | -8.968% | -9.188% | 4.123% | -8.803% | -11.33% | 1.126% | -15.29% | 13.52%
| Z-Score (Absolute) |  | 1.31 | 1.36 | 2.31 | 0.5 | 0 | 0.38 | 0.16 | 0.61 | 0.82 | 0.38 | 0.23 | 0.22 | 0.29 | 0.67 | 1.66 | 1.49 | 1.52 | 0.45 | 1.47 | 1.84 | 0.01 | 2.43 | 1.85
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 610 (Certified Value 48.4 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 1
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 2
| 1 | 1 | 49.9 | 53.1 | 49.869 | 47.5 | 48.5 | 49.3 | 48 | 49.6 | 48.4 | 50 | 48.33 | 49.9728 | 47.6 | 48.812 | 45.2 | 49.9919 | 48.5 | 50 | > 10 | 49.24 | 46 | 54.8 | 45 | 48.42 | 1
| 1 | 2 | 48.3 | 50.6 | 49.783 | 46.4 | 47.4 | 49.9 | 48.5 | 50.1 | 46.3 | 49.6 | 46.97 | 49.5745 | 47.2 | 47.349 | 45.7 | 50.2611 | 48.4 | 51 | > 10 | 51.14 | 46 | 56.3 | 45 | 48.02 | 27
| 1 | 3 | 48.3 | 52.5 | 49.459 | 44.7 | 48.4 | 49.7 | 48.3 | 50 | 48.1 | 50.8 | 47.58 | 48.5888 | 46.6 | 48.285 | 44.6 | 50.3504 | 49 | 50 | > 10 | 50.88 | 46 | 56.7 | 44 | 48.62 | 16
| 1 | 4 | 50.2 | 52.4 | 49.823 | 45 | 47.3 | 49.5 | 49 | 49.1 | 48.3 | 49.6 | 47.58 | 50.0375 | 47.5 | 48.521 | 44.1 | 49.5779 | 47.6 | 50 | > 10 | 50.35 | 45 | 55.8 | 44 | 47.28 | 48.4286
| 1 | 5 | 50.7 | 53 | 49.726 | 43.1 | 48 | 48.7 | 47.9 | 48.4 | 50.3 | 50.1 | 47.61 | 49.064 | 47.3 | 48.865 | 46 | 49.7482 | 48.5 | 50 | > 10 | 45.05 | 46 | 55.5 | 44 | 47.82 | 69
| 1 | 6 | 49.9 | 52.3 | 49.546 | 45.5 | 46.8 | 49.3 | 49.2 | 46.6 | 49.7 | 50.8 | 46.74 | 49.0067 | 47.8 | 47.887 | 45.3 | 50.0433 | 48.3 | 50 | > 10 | 50.79 | 46 | 55.9 | 44 | 48.42
| Mean |  | 49.55 | 52.3167 | 49.701 | 45.3667 | 47.7333 | 49.4 | 48.4833 | 48.9667 | 48.5167 | 50.15 | 47.4683 | 49.3741 | 47.3333 | 48.2865 | 45.15 | 49.9955 | 48.3833 | 50.1667 |  | 49.575 | 45.8333 | 55.8333 | 44.3333 | 48.0967
| Median |  | 49.9 | 52.45 | 49.7545 | 45.25 | 47.7 | 49.4 | 48.4 | 49.35 | 48.35 | 50.05 | 47.58 | 49.3193 | 47.4 | 48.403 | 45.25 | 50.0176 | 48.45 | 50 |  | 50.57 | 46 | 55.85 | 44 | 48.22
| Std Dev. |  | 1.0114 | 0.90203 | 0.16312 | 1.5069 | 0.67429 | 0.41473 | 0.52694 | 1.3186 | 1.3949 | 0.54314 | 0.55969 | 0.58078 | 0.41793 | 0.58383 | 0.70071 | 0.29476 | 0.4535 | 0.40825 |  | 2.3159 | 0.40825 | 0.65625 | 0.5164 | 0.49645
| Rel.Std.Dev. |  | 2.041% | 1.724% | 0.3282% | 3.322% | 1.413% | 0.8395% | 1.087% | 2.693% | 2.875% | 1.083% | 1.179% | 1.176% | 0.883% | 1.209% | 1.552% | 0.5896% | 0.9373% | 0.8138% |  | 4.672% | 0.8907% | 1.175% | 1.165% | 1.032%
| PDM3 |  | 2.316% | 8.029% | 2.627% | -6.323% | -1.436% | 2.006% | 0.1131% | 1.111% | 0.1819% | 3.555% | -1.983% | 1.952% | -2.262% | -0.2934% | -6.77% | 3.235% | -0.09342% | 3.589% |  | 2.367% | -5.359% | 15.29% | -8.456% | -0.6854%
| Z-Score (Absolute) |  | 0.66 | 2.42 | 0.75 | 2.01 | 0.5 | 0.56 | 0.02 | 0.29 | 0 | 1.04 | 0.67 | 0.55 | 0.75 | 0.15 | 2.15 | 0.94 | 0.08 | 1.05 | NA | 0.67 | 1.71 | 4.66 | 2.67 | 0.27
| Analytical results for Al in OREAS 610 (Certified Value 0.847 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 1.1608 | 0.93 | 1.37 | 0.85 | 0.92 | 0.87 | 0.88 | 0.89 | 1.19 | 0.913 | 1.1787 | 0.81 | 0.91 | 0.68 | 0.91028 | 0.75 | 0.855 | 0.84 | 0.87 | 0.86 | 0.85 | 0.83 | 0.8329 |  | 1
| 1 | 2 | 1.1393 | 0.96 | 1.4 | 0.86 | 0.91 | 0.88 | 0.9 | 0.87 | 1.23 | 0.894 | 1.1936 | 0.71 | 0.86 | 0.67 | 0.91326 | 0.76 | 0.864 | 0.84 | 0.89 | 0.86 | 0.84 | 0.83 | 0.7448 |  | #N/A
| 1 | 3 | 1.1246 | 0.95 | 1.32 | 0.85 | 0.91 | 0.87 | 0.88 | 0.88 | 1.24 | 0.773 | 1.1893 | 0.72 | 0.89 | 0.66 | 0.94513 | 0.76 | 0.868 | 0.84 | 0.87 | 0.88 | 0.84 | 0.82 | 0.7567 |  | 16
| 1 | 4 | 1.1273 | 0.94 | 1.3 | 0.84 | 0.89 | 0.86 | 0.87 | 0.9 | 1.16 | 0.764 | 1.1934 | 0.73 | 0.9 | 0.67 | 0.85443 | 0.75 | 0.868 | 0.84 | 0.86 | 0.87 | 0.78 | 0.81 | 0.7686 |  | 0.84724
| 1 | 5 | 1.1473 | 0.96 | 1.28 | 0.85 | 0.9 | 0.87 | 0.84 | 0.92 | 1.2 | 0.753 | 1.1862 | 0.73 | 0.9 | 0.64 | 0.83888 | 0.76 | 0.864 | 0.83 | 0.78 | 0.87 | 0.77 | 0.82 | 0.7805 |  | 70
| 1 | 6 | 1.1414 | 0.95 | 1.32 | 0.83 | 0.89 | 0.87 | 0.85 | 0.89 | 1.21 | 0.852 | 1.1799 | 0.74 | 0.89 | 0.61 | 0.82406 | 0.75 | 0.86 | 0.84 | 0.84 | 0.84 | 0.77 | 0.83 | 0.7924
| Mean |  | 1.1401 | 0.94833 | 1.3317 | 0.84667 | 0.90333 | 0.87 | 0.87 | 0.89167 | 1.205 | 0.82483 | 1.1868 | 0.74 | 0.89167 | 0.655 | 0.88101 | 0.755 | 0.86317 | 0.83833 | 0.85167 | 0.86333 | 0.80833 | 0.82333 | 0.77932
| Median |  | 1.1403 | 0.95 | 1.32 | 0.85 | 0.905 | 0.87 | 0.875 | 0.89 | 1.205 | 0.8125 | 1.1877 | 0.73 | 0.895 | 0.665 | 0.88235 | 0.755 | 0.864 | 0.84 | 0.865 | 0.865 | 0.81 | 0.825 | 0.77455
| Std Dev. |  | 0.013317 | 0.01169 | 0.044907 | 0.010328 | 0.012111 | 0.0063246 | 0.021909 | 0.017224 | 0.02881 | 0.070488 | 0.0064777 | 0.035777 | 0.017224 | 0.025884 | 0.048441 | 0.0054772 | 0.0049967 | 0.0040825 | 0.038687 | 0.013663 | 0.038687 | 0.008165 | 0.031182
| Rel.Std.Dev. |  | 1.168% | 1.233% | 3.372% | 1.22% | 1.341% | 0.727% | 2.518% | 1.932% | 2.391% | 8.546% | 0.5458% | 4.835% | 1.932% | 3.952% | 5.498% | 0.7255% | 0.5789% | 0.487% | 4.542% | 1.583% | 4.786% | 0.9917% | 4.001%
| PDM3 |  | 34.57% | 11.93% | 57.18% | -0.06769% | 6.621% | 2.686% | 2.686% | 5.244% | 42.23% | -2.645% | 40.08% | -12.66% | 5.244% | -22.69% | 3.985% | -10.89% | 1.88% | -1.051% | 0.5225% | 1.899% | -4.592% | -2.822% | -8.017%
| Z-Score (Absolute) |  | 4.67 | 1.43 | 7.9 | 0.28 | 0.67 | 0.11 | 0.11 | 0.48 | 5.76 | 0.65 | 5.45 | 2.08 | 0.48 | 3.51 | 0.3 | 1.83 | 0 | 0.42 | 0.2 | 0 | 0.93 | 0.67 | 1.42
| Analytical results for As in OREAS 610 (Certified Value 2807 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 0
| 1 | 1 | 2350 | 2591 | 2989.6 | 2820 | 2920 | 3080 | 2930 | 2920 | 3040 | 2570 | 2874.93 | 2810.7954 | 2818 | 2792 | 2820 | 2939.1732 | 2879 | 2928.945 | 2619 | 2503 | 2641 | > 1000 | 2918 | 2768 | 1
| 1 | 2 | 2210 | 2517 | 3043.7 | 2800 | 2900 | 3080 | 2910 | 2950 | 2900 | 2600 | 2836.25 | 2871.0159 | 2788 | 2700.8 | 2830 | 2958.9197 | 2862 | 2899.656 | 2606 | 2483 | 2648 | > 1000 | 2901 | 2757 | 17
| 1 | 3 | 2230 | 2481 | 3031 | 2630 | 2880 | 3060 | 2880 | 2970 | 2970 | 2660 | 2750.51 | 2735.1998 | 2773 | 2756.5 | 2880 | 2961.8942 | 2901 | 2890.419 | 2673 | 2539 | 2674 | > 1000 | 2889 | 2762 | 16
| 1 | 4 | 2420 | 2598 | 2972.3 | 2660 | 2850 | 3060 | 2910 | 2910 | 2990 | 2680 | 2754.66 | 2855.5406 | 2807 | 2755.5 | 2850 | 2912.0149 | 2829 | 2896.816 | 2573 | 2503 | 2589 | > 1000 | 2873 | 2755 | 2806.8049
| 1 | 5 | 2340 | 2561 | 3011.5 | 2560 | 2920 | 3050 | 2860 | 2840 | 3130 | 2610 | 2746.57 | 2778.5084 | 2825 | 2792.5 | 2790 | 2888.5073 | 2886 | 2894.107 | 2642 | 2320 | 2670 | > 1000 | 2844 | 2760 | 71
| 1 | 6 | 2360 | 2598 | 3030.2 | 2690 | 2850 | 3040 | 2890 | 2840 | 3050 | 2620 | 2803.54 | 2764.194 | 2838 | 2733.5 | 2810 | 2935.7698 | 2851 | 2898.01 | 2590 | 2483 | 2636 | > 1000 | 2886 | 2753
| Mean |  | 2318.3333 | 2557.6667 | 3013.05 | 2693.3333 | 2886.6667 | 3061.6667 | 2896.6667 | 2905 | 3013.3333 | 2623.3333 | 2794.41 | 2802.5423 | 2808.1667 | 2755.1333 | 2830 | 2932.7132 | 2868 | 2901.3255 | 2617.1667 | 2471.8333 | 2643 |  | 2885.1667 | 2759.1667
| Median |  | 2345 | 2576 | 3020.85 | 2675 | 2890 | 3060 | 2900 | 2915 | 3015 | 2615 | 2779.1 | 2794.6519 | 2812.5 | 2756 | 2825 | 2937.4715 | 2870.5 | 2897.413 | 2612.5 | 2493 | 2644.5 |  | 2887.5 | 2758.5
| Std Dev. |  | 81.3429 | 48.8003 | 27.454 | 100.3328 | 32.0416 | 16.0208 | 25.0333 | 54.6809 | 78.6554 | 40.332 | 53.1288 | 53.1955 | 24.1447 | 35.1498 | 31.6228 | 28.185 | 26 | 13.9141 | 36.1962 | 77.1451 | 30.6072 |  | 25.2461 | 5.4191
| Rel.Std.Dev. |  | 3.509% | 1.908% | 0.9112% | 3.725% | 1.11% | 0.5233% | 0.8642% | 1.882% | 2.61% | 1.537% | 1.901% | 1.898% | 0.8598% | 1.276% | 1.117% | 0.9611% | 0.9066% | 0.4796% | 1.383% | 3.121% | 1.158% |  | 0.875% | 0.1964%
| PDM3 |  | -17.4% | -8.876% | 7.348% | -4.043% | 2.845% | 9.08% | 3.202% | 3.498% | 7.358% | -6.537% | -0.4416% | -0.1519% | 0.04852% | -1.841% | 0.8264% | 4.486% | 2.18% | 3.368% | -6.756% | -11.93% | -5.836% |  | 2.792% | -1.697%
| Z-Score (Absolute) |  | 3.41 | 1.74 | 1.43 | 0.8 | 0.55 | 1.77 | 0.62 | 0.67 | 1.43 | 1.29 | 0.1 | 0.04 | 0 | 0.37 | 0.15 | 0.87 | 0.42 | 0.65 | 1.33 | 2.34 | 1.15 | NA | 0.54 | 0.34
| Analytical results for B in OREAS 610 (Indicative Value < 10 ppm) | [25 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.20 | 2.24 | [8 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | [8 empty] | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | [8 empty] | 1
| 1 | 1 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 1
| 1 | 2 | < 10 | 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 8
| 1 | 3 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 16
| 1 | 4 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | < 10
| 1 | 5 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 14
| 1 | 6 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15
| Mean |  | 2.03: 1; 2.24: 15
| Median |  | 2.03: 1; 2.24: 15
| Std Dev. |  | 2.24: 0
Rel.Std.Dev. | bar: 2.24
PDM3
Z-Score (Absolute) | bar: 2.02 to 2.24
| Analytical results for Ba in OREAS 610 (Indicative Value 84 ppm) | [25 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  | 0
| 1 | 1 | 340 | 26.1 | 19.5 | 30 | 30 | 30 | 30 | 30 | 255 | 14.46 | 87.317 | 241 | 24.9 | 141.8558 | 308 | 124.32 | 166 | 111 | 140 | 72 | 24 | 2526 |  |  | 1
| 1 | 2 | 315 | 27.3 | 14.4 | 40 | 40 | 30 | 30 | 30 | 257 | 14.93 | 86.2009 | 188 | 26.3 | 122.2055 | 313 | 137.341 | 167 | 111 | 145 | 77 | 24 | 2517 |  |  | 5
| 1 | 3 | 367 | 26.6 | 15.4 | 40 | 30 | 30 | 30 | 30 | 259 | 14.14 | 74.6012 | 183 | 24.7 | 150.0214 | 307 | 143.193 | 158 | 126 | 145 | 69 | 24 | 2508 |  |  | 16
| 1 | 4 | 464 | 25.1 | 18 | 40 | 30 | 30 | 30 | 30 | 249 | 12.82 | 72.2197 | 225 | 25 | 145.3821 | 298 | 150.254 | 170 | 111 | 142 | 73 | 23 | 2523 |  |  | 83.9675
| 1 | 5 | 370 | 28.3 | 15.3 | 30 | 40 | 30 | 30 | 30 | 261 | 14.5 | 81.2618 | 208 | 39.4 | 163.0945 | 298 | 146.262 | 161 | 115 | 146 | 74 | 23 | 2580 |  |  | 15
| 1 | 6 | 398 | 28.3 | 17.7 | 30 | 30 | 40 | 50 | 30 | 260 | 13.45 | 71.4219 | 199 | 24.2 | 149.5222 | 304 | 148.095 | 158 | 115 | 147 | 69 | 23 | 2581
| Mean |  | 375.6667 | 26.95 | 16.7167 | 35 | 33.3333 | 31.6667 | 33.3333 | 30 | 256.8333 | 14.05 | 78.8371 | 207.3333 | 27.4167 | 145.3469 | 304.6667 | 141.5775 | 163.3333 | 114.8333 | 144.1667 | 72.3333 | 23.5 | 2539.1667
| Median |  | 368.5 | 26.95 | 16.55 | 35 | 30 | 30 | 30 | 30 | 258 | 14.3 | 77.9315 | 203.5 | 24.95 | 147.4521 | 305.5 | 144.7275 | 163.5 | 113 | 145 | 72.5 | 23.5 | 2524.5
| Std Dev. |  | 51.694 | 1.2677 | 1.9732 | 5.4772 | 5.164 | 4.0825 | 8.165 | 0 | 4.4008 | 0.77795 | 7.0531 | 22.277 | 5.9118 | 13.4311 | 5.9217 | 9.5717 | 5.0465 | 5.8109 | 2.6394 | 3.0768 | 0.54772 | 32.6032
| Rel.Std.Dev. |  | 13.76% | 4.704% | 11.8% | 15.65% | 15.49% | 12.89% | 24.49% | 0% | 1.713% | 5.537% | 8.946% | 10.74% | 21.56% | 9.241% | 1.944% | 6.761% | 3.09% | 5.06% | 1.831% | 4.254% | 2.331% | 1.284%
| PDM3 |  | 347.4% | -67.9% | -80.09% | -58.32% | -60.3% | -62.29% | -60.3% | -64.27% | 205.9% | -83.27% | -6.11% | 146.9% | -67.35% | 73.1% | 262.8% | 68.61% | 94.52% | 36.76% | 71.69% | -13.86% | -72.01% | 2924%
| Z-Score (Absolute) |  | 3.65 | 0.59 | 0.72 | 0.49 | 0.51 | 0.53 | 0.51 | 0.55 | 2.2 | 0.75 | 0.04 | 1.6 | 0.59 | 0.85 | 2.78 | 0.8 | 1.07 | 0.48 | 0.83 | 0.04 | 0.63 | 29.94
| Analytical results for Be in OREAS 610 (Certified Value 0.29 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 01g | [5 empty] | 3
| 1 | 1 | < 0.5 | 0.4 | 0.4 | 0.24 | 0.31 | 0.28 | 0.3 | 0.3 | < 0.5 | 0.31 | < 0.5 | 0.31 | < 0.5 | < 0.5 | 0.4 | 0.3 | < 0.5 | < 1 | < 1 | [5 empty] | 1
| 1 | 2 | < 0.5 | 0.2 | 0.4 | 0.26 | 0.26 | 0.28 | 0.29 | 0.32 | < 0.5 | 0.28 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.3 | 0.3 | < 0.5 | < 1 | < 1 | [5 empty] | 30
| 1 | 3 | < 0.5 | 0.3 | 0.3 | 0.36 | 0.31 | 0.25 | 0.28 | 0.31 | < 0.5 | 0.3 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.3 | 0.3 | < 0.5 | < 1 | < 1 | [5 empty] | 16
| 1 | 4 | < 0.5 | 0.3 | 0.3 | 0.29 | 0.27 | 0.29 | 0.28 | 0.29 | < 0.5 | 0.31 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.3 | 0.3 | < 0.5 | < 1 | < 1 | [5 empty] | 0.28972
| 1 | 5 | < 0.5 | 0.3 | 0.4 | 0.31 | 0.31 | 0.29 | 0.25 | 0.33 | < 0.5 | 0.28 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.3 | 0.2 | < 0.5 | < 1 | < 1 | [5 empty] | 72
| 1 | 6 | < 0.5 | 0.3 | 0.3 | 0.29 | 0.26 | 0.3 | 0.34 | 0.34 | < 0.5 | 0.29 | < 0.5 | 0.28 | < 0.5 | < 0.5 | 0.3 | 0.3 | < 0.5 | < 1 | < 1
| Mean |  | 2.03: 0.3; 2.04: 0.35; 2.05: 0.29167; 2.06: 0.28667; 2.07: 0.28167; 2.08: 0.29; 2.09: 0.315; 2.11: 0.295; 2.15: 0.29333; 2.20: 0.31667; 2.21: 0.28333
| Median |  | 2.03: 0.3; 2.04: 0.35; 2.05: 0.29; 2.06: 0.29; 2.07: 0.285; 2.08: 0.285; 2.09: 0.315; 2.11: 0.295; 2.15: 0.29; 2.20: 0.3; 2.21: 0.3
| Std Dev. |  | 2.03: 0.063246; 2.04: 0.054772; 2.05: 0.041673; 2.06: 0.02582; 2.07: 0.017224; 2.08: 0.029665; 2.09: 0.018708; 2.11: 0.013784; 2.15: 0.010328; 2.20: 0.040825; 2.21: 0.040825
| Rel.Std.Dev. |  | 2.03: 21.08%; 2.04: 15.65%; 2.05: 14.29%; 2.06: 9.007%; 2.07: 6.115%; 2.08: 10.23%; 2.09: 5.939%; 2.11: 4.673%; 2.15: 3.521%; 2.20: 12.89%; 2.21: 14.41%
| PDM3 |  | 2.03: 3.547%; 2.04: 20.81%; 2.05: 0.6711%; 2.06: -1.055%; 2.07: -2.78%; 2.08: 0.09588%; 2.09: 8.725%; 2.11: 1.822%; 2.15: 1.246%; 2.20: 9.3%; 2.21: -2.205%
| Z-Score (Absolute) |  | 0.67 | NA | NA | 0.09 | 0 | 0.09 | 0.06 | 0.52 | 0.67 | 0.15 | 0.67 | 0.12 | 0.67 | 0.67 | NA | NA | 0.67 | 3.92 | 3.92
Results from laboratories 3, 4, 20 & 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 610 (Certified Value 220 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 198.57 | 225.95 | 198 | 220 | 233 | 224 | 222 | 227 | 231 | 222.34 | 197.2411 | 240 | 207.64 | 233 | 225.0891 | 204.6 | 225.494 | 215.89 | 210 | 213 | 225.7 | 323 | 136 |  | 1
| 1 | 2 | 194.91 | 226.12 | 194 | 223 | 233 | 219 | 222 | 216 | 233 | 221.4 | 203.5584 | 239 | 201.24 | 234 | 209.6151 | 206.39 | 222.269 | 220.17 | 219 | 214 | 219.6 | 325 | 137.4 |  | 31
| 1 | 3 | 199.92 | 228.55 | 189 | 219 | 232 | 218 | 226 | 223 | 237 | 225.35 | 198.3745 | 235 | 205.16 | 235 | 219.2207 | 206.88 | 221.184 | 215.96 | 213 | 215 | 224.3 | 318 | 136.8 |  | 16
| 1 | 4 | 197.21 | 226.58 | 190 | 216 | 232 | 222 | 222 | 224 | 234 | 222.55 | 202.3414 | 239 | 204.99 | 230 | 227.0083 | 198.8 | 223.772 | 216.99 | 208 | 206 | 222.4 | 318 | 138.2 |  | 219.8074
| 1 | 5 | 202.72 | 222.75 | 183 | 224 | 234 | 219 | 213 | 237 | 235 | 220.6 | 196.2131 | 240 | 208.74 | 223 | 232.0304 | 205.09 | 224.566 | 220.96 | 194 | 213 | 224.8 | 319 | 140.6 |  | 73
| 1 | 6 | 200.48 | 221.88 | 191 | 217 | 233 | 221 | 216 | 229 | 231 | 217.06 | 199.839 | 242 | 204.34 | 234 | 218.6827 | 206.82 | 225.444 | 217.69 | 213 | 216 | 225.7 | 316 | 140
| Mean |  | 198.9683 | 225.305 | 190.8333 | 219.8333 | 232.8333 | 220.5 | 220.1667 | 226 | 233.5 | 221.55 | 199.5946 | 239.1667 | 205.3517 | 231.5 | 221.941 | 204.7633 | 223.7882 | 217.9433 | 209.5 | 212.8333 | 223.75 | 319.8333 | 138.1667
| Median |  | 199.245 | 226.035 | 190.5 | 219.5 | 233 | 220 | 222 | 225.5 | 233.5 | 221.87 | 199.1067 | 239.5 | 205.075 | 233.5 | 222.1549 | 205.74 | 224.169 | 217.34 | 211.5 | 213.5 | 224.55 | 318.5 | 137.8
| Std Dev. |  | 2.7198 | 2.5098 | 5.0365 | 3.1885 | 0.75277 | 2.2583 | 4.7504 | 6.9857 | 2.3452 | 2.7258 | 2.8906 | 2.3166 | 2.6387 | 4.5056 | 7.84 | 3.0676 | 1.7519 | 2.1532 | 8.4558 | 3.5449 | 2.3688 | 3.4303 | 1.8129
| Rel.Std.Dev. |  | 1.367% | 1.114% | 2.639% | 1.45% | 0.3233% | 1.024% | 2.158% | 3.091% | 1.004% | 1.23% | 1.448% | 0.9686% | 1.285% | 1.946% | 3.532% | 1.498% | 0.7829% | 0.988% | 4.036% | 1.666% | 1.059% | 1.073% | 1.312%
| PDM3 |  | -9.481% | 2.501% | -13.18% | 0.0118% | 5.926% | 0.3151% | 0.1634% | 2.817% | 6.229% | 0.7928% | -9.196% | 8.807% | -6.577% | 5.319% | 0.9707% | -6.844% | 1.811% | -0.848% | -4.689% | -3.173% | 1.794% | 45.51% | -37.14%
| Z-Score (Absolute) |  | 1.32 | 0.29 | 1.82 | 0.04 | 0.76 | 0 | 0.02 | 0.34 | 0.8 | 0.06 | 1.28 | 1.14 | 0.93 | 0.67 | 0.09 | 0.96 | 0.2 | 0.16 | 0.67 | 0.47 | 0.2 | 6.09 | 5.05
| Analytical results for Ca in OREAS 610 (Certified Value 0.12 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.13 | 0.12 | 0.13 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.117 | 0.12767 | 0.11 | 0.12 | 0.13 | 0.12391 | 0.11 | 0.119 | 0.13 | 0.12 | 0.11 | 0.13 | 0.11 | 0.3367 |  | 1
| 1 | 2 | 0.14 | 0.12 | 0.13 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.112 | 0.1255 | 0.11 | 0.12 | 0.13 | 0.12309 | 0.11 | 0.119 | 0.12 | 0.12 | 0.11 | 0.13 | 0.11 | 0.362 |  | #N/A
| 1 | 3 | 0.14 | 0.12 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.098 | 0.1189 | 0.11 | 0.12 | 0.12 | 0.12559 | 0.11 | 0.119 | 0.13 | 0.12 | 0.12 | 0.13 | 0.11 | 0.3657 |  | 16
| 1 | 4 | 0.13 | 0.12 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.097 | 0.12102 | 0.11 | 0.12 | 0.12 | 0.12288 | 0.11 | 0.12 | 0.12 | 0.12 | 0.12 | 0.13 | 0.11 | 0.3752 |  | 0.11989
| 1 | 5 | 0.14 | 0.12 | 0.11 | 0.12 | 0.13 | 0.12 | 0.12 | 0.13 | 0.12 | 0.097 | 0.11732 | 0.11 | 0.12 | 0.12 | 0.12358 | 0.11 | 0.119 | 0.13 | 0.11 | 0.12 | 0.13 | 0.11 | 0.3547 |  | 74
| 1 | 6 | 0.14 | 0.12 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.109 | 0.12004 | 0.11 | 0.12 | 0.11 | 0.12255 | 0.11 | 0.119 | 0.12 | 0.12 | 0.11 | 0.13 | 0.11 | 0.3342
| Mean |  | 0.13667 | 0.12 | 0.12167 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12167 | 0.12 | 0.105 | 0.12174 | 0.11 | 0.12 | 0.12167 | 0.1236 | 0.11 | 0.11917 | 0.125 | 0.11833 | 0.115 | 0.13 | 0.11 | 0.35475
| Median |  | 0.14 | 0.12 | 0.12 | 0.12 | 0.13 | 0.12 | 0.12 | 0.12 | 0.12 | 0.1035 | 0.12053 | 0.11 | 0.12 | 0.12 | 0.12333 | 0.11 | 0.119 | 0.125 | 0.12 | 0.115 | 0.13 | 0.11 | 0.35835
| Std Dev. |  | 0.005164 | 0 | 0.0075277 | 0 | 0 | 0 | 0 | 0.0040825 | 0 | 0.0087864 | 0.0040089 | 0 | 0 | 0.0075277 | 0.0010907 | 0 | 0.00040825 | 0.0054772 | 0.0040825 | 0.0054772 | 0 | 0 | 0.016363
| Rel.Std.Dev. |  | 3.779% | 0% | 6.187% | 0% | 0% | 0% | 0% | 3.355% | 0% | 8.368% | 3.293% | 0% | 0% | 6.187% | 0.8824% | 0% | 0.3426% | 4.382% | 3.45% | 4.763% | 0% | 0% | 4.613%
| PDM3 |  | 13.99% | 0.08999% | 1.48% | 0.08999% | 8.431% | 0.08999% | 0.08999% | 1.48% | 0.08999% | -12.42% | 1.543% | -8.251% | 0.08999% | 1.48% | 3.093% | -8.251% | -0.6051% | 4.26% | -1.3% | -4.08% | 8.431% | -8.251% | 195.9%
| Z-Score (Absolute) |  | 6.45 | 0 | 0.65 | 0 | 3.87 | 0 | 0 | 0.65 | 0 | 5.81 | 0.67 | 3.87 | 0 | 0.65 | 1.39 | 3.87 | 0.32 | 1.94 | 0.65 | 1.94 | 3.87 | 3.87 | 90.88
| Analytical results for Cd in OREAS 610 (Certified Value 12.3 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 2
| 1 | 1 | 10.49 | 13.15 | 11.4 | 12 | 12.15 | 11.7 | 12.55 | 12.35 | 10.2 | 11.819 | 12.6216 | 11.8 | 13.32 | 12.7 | 12.3158 | 10.51 | 12.182 | 12.6 | 11.79 | 11 | 14.6 | 27 | 12.45 |  | 1
| 1 | 2 | 10.47 | 13.34 | 11 | 11.85 | 12.3 | 11.9 | 12.2 | 11.9 | 10 | 11.672 | 12.9246 | 11.7 | 12.47 | 12.5 | 12.1568 | 10.68 | 12.055 | 12.67 | 11.6 | 12 | 14.4 | 28 | 12.7 |  | 32
| 1 | 3 | 10.4 | 13.2 | 10.7 | 12.75 | 12.15 | 11.9 | 12.35 | 12.35 | 10.6 | 11.976 | 12.2321 | 11.6 | 12.4 | 12.4 | 12.2898 | 10.51 | 12.121 | 12.68 | 11.33 | 12 | 14.6 | 28 | 12.55 |  | 16
| 1 | 4 | 10.43 | 13.33 | 10.8 | 12.55 | 11.95 | 12.35 | 12 | 12.4 | 10.6 | 11.904 | 12.966 | 12 | 13.38 | 12.8 | 12.2972 | 10.27 | 12.194 | 12.23 | 11.47 | 11 | 14.6 | 28 | 12.2 |  | 12.253
| 1 | 5 | 10.48 | 13.24 | 10.5 | 11.85 | 11.8 | 11.9 | 11.7 | 12.95 | 10.4 | 11.865 | 12.5342 | 12.1 | 12.64 | 12.8 | 12.3292 | 10.58 | 12.145 | 12.64 | 10.47 | 12 | 14.5 | 28 | 13.45 |  | 75
| 1 | 6 | 10.82 | 12.79 | 10.9 | 12.1 | 12.3 | 12.2 | 13.05 | 12.85 | 10.2 | 11.705 | 12.3181 | 12.3 | 12.99 | 12.8 | 12.1949 | 10.68 | 12.114 | 12.53 | 11.49 | 12 | 14.6 | 28 | 12.7
| Mean |  | 10.515 | 13.175 | 10.8833 | 12.1833 | 12.1083 | 11.9917 | 12.3083 | 12.4667 | 10.3333 | 11.8235 | 12.5994 | 11.9167 | 12.8667 | 12.6667 | 12.2639 | 10.5383 | 12.1352 | 12.5583 | 11.3583 | 11.6667 | 14.55 | 27.8333 | 12.675
| Median |  | 10.475 | 13.22 | 10.85 | 12.05 | 12.15 | 11.9 | 12.275 | 12.375 | 10.3 | 11.842 | 12.5779 | 11.9 | 12.815 | 12.75 | 12.2935 | 10.545 | 12.133 | 12.62 | 11.48 | 12 | 14.6 | 28 | 12.625
| Std Dev. |  | 0.1532 | 0.20246 | 0.30605 | 0.37903 | 0.19854 | 0.23752 | 0.46628 | 0.38297 | 0.24221 | 0.11702 | 0.30296 | 0.26394 | 0.42679 | 0.17512 | 0.070652 | 0.15198 | 0.050649 | 0.16987 | 0.46141 | 0.5164 | 0.083666 | 0.40825 | 0.42279
| Rel.Std.Dev. |  | 1.457% | 1.537% | 2.812% | 3.111% | 1.64% | 1.981% | 3.788% | 3.072% | 2.344% | 0.9898% | 2.405% | 2.215% | 3.317% | 1.383% | 0.5761% | 1.442% | 0.4174% | 1.353% | 4.062% | 4.426% | 0.575% | 1.467% | 3.336%
| PDM3 |  | -14.18% | 7.525% | -11.18% | -0.5684% | -1.181% | -2.133% | 0.4517% | 1.744% | -15.67% | -3.505% | 2.827% | -2.745% | 5.008% | 3.376% | 0.08936% | -13.99% | -0.9615% | 2.492% | -7.301% | -4.785% | 18.75% | 127.2% | 3.444%
| Z-Score (Absolute) |  | 2.7 | 1.61 | 2.11 | 0 | 0.12 | 0.31 | 0.2 | 0.46 | 3 | 0.58 | 0.67 | 0.43 | 1.11 | 0.78 | 0.13 | 2.67 | 0.08 | 0.61 | 1.34 | NA | 3.84 | NA | 0.8
Results from laboratories 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 610 (Certified Value 13.7 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 1
|  |  | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 2
| 1 | 1 | 13.2 | 18 | 14.2 | 13.5 | 13.65 | 14.6 | 13.85 | 23.2 | 14.802 | 19.7613 | 11.7 | 12.9339 | 13.7 | 15.31 | 14.56 | 1 | [8 empty] | 1
| 1 | 2 | 13.9 | 17.5 | 14.3 | 13.6 | 14.35 | 14.7 | 13.85 | 23.7 | 14.756 | 20.1983 | 12 | 11.9383 | 13.4 | 14.99 | 14.88 | 1 | [8 empty] | 33
| 1 | 3 | 13.8 | 17 | 13.9 | 13.3 | 13.85 | 13.85 | 13.9 | 24.1 | 15.345 | 19.5346 | 12 | 14.2741 | 13.4 | 14.99 | 14.3 | 1 | [8 empty] | 16
| 1 | 4 | 13.7 | 17.4 | 14.05 | 12.9 | 14.4 | 13.7 | 13.9 | 22.5 | 14.816 | 20.193 | 11.6 | 11.373 | 13.7 | 15.22 | 14.11 | 1 | [8 empty] | 13.7282
| 1 | 5 | 13.9 | 16.1 | 13.5 | 13.25 | 14.1 | 13.55 | 14.35 | 23.3 | 14.945 | 19.6962 | 11.5 | 11.7381 | 13.8 | 15.23 | 13.12 | 1 | [8 empty] | 76
| 1 | 6 | 13.1 | 16.8 | 13.85 | 13.05 | 14.5 | 15.1 | 14.35 | 23.2 | 14.665 | 19.9316 | 11.7 | 10.5623 | 12.7 | 15.18 | 14.19 | < 1
| Mean |  | 13.6 | 17.1333 | 13.9667 | 13.2667 | 14.1417 | 14.25 | 14.0333 | 23.3333 | 14.8882 | 19.8858 | 11.75 | 12.1366 | 13.45 | 15.1533 | 14.1933 | 1
| Median |  | 13.75 | 17.2 | 13.975 | 13.275 | 14.225 | 14.225 | 13.9 | 23.25 | 14.809 | 19.8465 | 11.7 | 11.8382 | 13.55 | 15.2 | 14.245 | 1
| Std Dev. |  | 0.35777 | 0.65625 | 0.28577 | 0.26394 | 0.33677 | 0.63246 | 0.24631 | 0.53914 | 0.24159 | 0.27161 | 0.20736 | 1.3011 | 0.40373 | 0.13337 | 0.59604 | 0
| Rel.Std.Dev. |  | 2.631% | 3.83% | 2.046% | 1.99% | 2.381% | 4.438% | 1.755% | 2.311% | 1.623% | 1.366% | 1.765% | 10.72% | 3.002% | 0.8801% | 4.199% | 0%
| PDM3 |  | -0.9339% | 24.8% | 1.737% | -3.362% | 3.012% | 3.801% | 2.223% | 69.97% | 8.45% | 44.85% | -14.41% | -11.59% | -2.026% | 10.38% | 3.388% | -92.72%
| Z-Score (Absolute) |  | 0.49 | 2.7 | 0.16 | 0.79 | 0 | 0.1 | 0.1 | 8.3 | 0.67 | 5.19 | 2.16 | 1.81 | 0.62 | 0.91 | 0.05 | NA
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 610 (Certified Value 7.73 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 2
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 3
| 1 | 1 | 6.18 | 7.8 | 8.4 | 8.5 | 7.8 | 7.6 | 7.7 | 7.6 | 8.1 | 7.8 | 7.97 | 7.7135 | 8 | 7.1 | 8.3 | 7.7946 | 7.17 | 7.572 | 7.6 | 7.6 | 7 | 8.5 | 7 | 2.412 | 1
| 1 | 2 | 6.13 | 7.4 | 8.7 | 7.9 | 7.3 | 7.8 | 8 | 7.6 | 7.9 | 7.8 | 7.68 | 7.805 | 7 | 6.5 | 8.5 | 7.7523 | 7.24 | 7.43 | 7.2 | 7.5 | 7 | 8.4 | 7 | 2.03 | 21
| 1 | 3 | 6.14 | 7.4 | 8.4 | 7.5 | 7.7 | 8 | 7.8 | 7.3 | 8 | 8 | 7.42 | 7.4971 | 7 | 7 | 8.2 | 7.7646 | 7.17 | 7.576 | 7.4 | 7.6 | 7 | 8.7 | 7 | 2.648 | 16
| 1 | 4 | 6.29 | 7.9 | 8.6 | 7.9 | 7.4 | 7.5 | 8.1 | 7.2 | 8.2 | 7.7 | 7.5 | 7.8571 | 8 | 6.7 | 8.3 | 7.4721 | 7.15 | 7.546 | 7.4 | 7.4 | 7 | 8.5 | 7 | 2.734 | 7.7258
| 1 | 5 | 5.75 | 7.7 | 8.3 | 7.3 | 7.5 | 7.6 | 7.8 | 7.4 | 8.3 | 7.8 | 7.34 | 7.5443 | 8 | 6.7 | 8.3 | 7.7089 | 7.39 | 7.638 | 7.3 | 6.8 | 7 | 8.5 | 7 | 2.116 | 77
| 1 | 6 | 5.86 | 7.8 | 8.4 | 7.9 | 7.3 | 7.8 | 7.9 | 6.8 | 8.4 | 8 | 7.58 | 7.6444 | 8 | 6.8 | 8.3 | 7.5362 | 7.28 | 7.672 | 7.4 | 7.5 | 7 | 8.6 | 7 | 2.498
| Mean |  | 6.0583 | 7.6667 | 8.4667 | 7.8333 | 7.5 | 7.7167 | 7.8833 | 7.3167 | 8.15 | 7.85 | 7.5817 | 7.6769 | 7.6667 | 6.8 | 8.3167 | 7.6714 | 7.2333 | 7.5723 | 7.3833 | 7.4 | 7 | 8.5333 | 7 | 2.4063
| Median |  | 6.135 | 7.75 | 8.4 | 7.9 | 7.45 | 7.7 | 7.85 | 7.35 | 8.15 | 7.8 | 7.54 | 7.679 | 8 | 6.75 | 8.3 | 7.7306 | 7.205 | 7.574 | 7.4 | 7.5 | 7 | 8.5 | 7 | 2.455
| Std Dev. |  | 0.20721 | 0.21602 | 0.15055 | 0.41312 | 0.20976 | 0.18348 | 0.1472 | 0.29944 | 0.18708 | 0.12247 | 0.22436 | 0.14227 | 0.5164 | 0.21909 | 0.098319 | 0.13403 | 0.09136 | 0.083918 | 0.13292 | 0.30332 | 0 | 0.10328 | 0 | 0.28289
| Rel.Std.Dev. |  | 3.42% | 2.818% | 1.778% | 5.274% | 2.797% | 2.378% | 1.867% | 4.093% | 2.295% | 1.56% | 2.959% | 1.853% | 6.736% | 3.222% | 1.182% | 1.747% | 1.263% | 1.108% | 1.8% | 4.099% | 0% | 1.21% | 0% | 11.76%
| PDM3 |  | -21.58% | -0.7657% | 9.589% | 1.392% | -2.923% | -0.1185% | 2.039% | -5.296% | 5.49% | 1.607% | -1.866% | -0.6332% | -0.7657% | -11.98% | 7.648% | -0.704% | -6.375% | -1.987% | -4.433% | -4.217% | -9.395% | 10.45% | -9.395% | -68.85%
| Z-Score (Absolute) |  | 4.07 | 0 | 2.02 | 0.42 | 0.42 | 0.13 | 0.55 | 0.89 | 1.22 | 0.46 | 0.21 | 0.03 | NA | 2.19 | 1.64 | 0.01 | 1.1 | 0.24 | 0.72 | 0.67 | NA | 2.19 | NA | 13.3
Results from laboratories 13, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 610 (Certified Value 33.1 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g |  | 2
| 1 | 1 | 24.1 | 38 | 36.3 | 36 | 33 | 35 | 34 | 34 | 34 | 35 | 33 | 32.9202 | 31 | 35.3 | 30.6 | 26.9373 | 35 | 31.412 | 35 | 32 | 31 | 36 | 30 |  | 1
| 1 | 2 | 21.6 | 36 | 36.5 | 35 | 33 | 36 | 34 | 34 | 32 | 36 | 32.4 | 33.4871 | 30 | 33.1 | 32.7 | 26.5458 | 34 | 31.653 | 36 | 32 | 31 | 36 | 30 |  | 35
| 1 | 3 | 20.3 | 37 | 37.1 | 34 | 33 | 35 | 33 | 34 | 34 | 36 | 30.9 | 34.3622 | 30 | 33.6 | 31.8 | 27.2538 | 34 | 31.608 | 34 | 32 | 31 | 35 | 30 |  | 16
| 1 | 4 | 23.4 | 37 | 36.2 | 35 | 33 | 36 | 33 | 33 | 34 | 35 | 30.5 | 33.9483 | 30 | 34.1 | 31.4 | 25.7911 | 36 | 31.364 | 35 | 31 | 30 | 35 | 30 |  | 33.0658
| 1 | 5 | 23.9 | 38 | 36.5 | 33 | 33 | 35 | 33 | 32 | 35 | 36 | 30.8 | 32.5648 | 30 | 34.5 | 31.2 | 25.7964 | 34 | 30.952 | 35 | 29 | 31 | 35 | 30 |  | 78
| 1 | 6 | 19 | 37 | 35.9 | 34 | 32 | 35 | 33 | 33 | 34 | 36 | 31.9 | 32.7586 | 30 | 35.2 | 31.3 | 25.4198 | 34 | 31.107 | 34 | 31 | 30 | 35 | 30
| Mean |  | 22.05 | 37.1667 | 36.4167 | 34.5 | 32.8333 | 35.3333 | 33.3333 | 33.3333 | 33.8333 | 35.6667 | 31.5833 | 33.3402 | 30.1667 | 34.3 | 31.5 | 26.2907 | 34.5 | 31.3493 | 34.8333 | 31.1667 | 30.6667 | 35.3333 | 30
| Median |  | 22.5 | 37 | 36.4 | 34.5 | 33 | 35 | 33 | 33.5 | 34 | 36 | 31.4 | 33.2037 | 30 | 34.3 | 31.35 | 26.1711 | 34 | 31.388 | 35 | 31.5 | 31 | 35 | 30
| Std Dev. |  | 2.0983 | 0.75277 | 0.40208 | 1.0488 | 0.40825 | 0.5164 | 0.5164 | 0.8165 | 0.98319 | 0.5164 | 1.0028 | 0.71441 | 0.40825 | 0.87407 | 0.70427 | 0.72981 | 0.83666 | 0.27563 | 0.75277 | 1.169 | 0.5164 | 0.5164 | 0
| Rel.Std.Dev. |  | 9.516% | 2.025% | 1.104% | 3.04% | 1.243% | 1.462% | 1.549% | 2.449% | 2.906% | 1.448% | 3.175% | 2.143% | 1.353% | 2.548% | 2.236% | 2.776% | 2.425% | 0.8792% | 2.161% | 3.751% | 1.684% | 1.462% | 0%
| PDM3 |  | -33.31% | 12.4% | 10.13% | 4.338% | -0.7029% | 6.858% | 0.8092% | 0.8092% | 2.321% | 7.866% | -4.483% | 0.83% | -8.768% | 3.733% | -4.735% | -20.49% | 4.338% | -5.191% | 5.346% | -5.743% | -7.256% | 6.858% | -9.272%
| Z-Score (Absolute) |  | 3.83 | 1.3 | 1.05 | 0.4 | 0.17 | 0.68 | 0 | 0 | 0.17 | 0.79 | 0.59 | 0 | 1.08 | 0.33 | 0.62 | 2.39 | 0.4 | 0.67 | 0.51 | 0.74 | 0.91 | 0.68 | 1.13
| Analytical results for Cs in OREAS 610 (Certified Value 0.74 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 0.89 | 0.75 | 0.86 | 0.74 | 0.72 | 0.72 | 0.75 | 0.77 | 0.89 | 0.77 | 0.97108 | 0.68 | 0.7 | 0.76 | 0.71 | 0.8 | [8 empty] | 1
| 1 | 2 | 0.89 | 0.77 | 0.86 | 0.72 | 0.73 | 0.74 | 0.73 | 0.74 | 0.9 | 0.77 | 0.97223 | 0.68 | 0.71 | 0.74 | 0.72 | 0.8 | [8 empty] | 36
| 1 | 3 | 0.88 | 0.76 | 0.75 | 0.74 | 0.73 | 0.74 | 0.74 | 0.75 | 0.92 | 0.78 | 0.93228 | 0.69 | 0.7 | 0.75 | 0.72 | 0.8 | [8 empty] | 16
| 1 | 4 | 0.91 | 0.76 | 0.77 | 0.72 | 0.69 | 0.74 | 0.73 | 0.78 | 0.89 | 0.8 | 0.96489 | 0.7 | 0.7 | 0.74 | 0.7 | 0.7 | [8 empty] | 0.73729
| 1 | 5 | 0.89 | 0.78 | 0.8 | 0.72 | 0.71 | 0.72 | 0.7 | 0.8 | 0.89 | 0.76 | 0.93413 | 0.71 | 0.69 | 0.75 | 0.65 | 0.7 | [8 empty] | 79
| 1 | 6 | 0.92 | 0.77 | 0.81 | 0.72 | 0.71 | 0.75 | 0.76 | 0.79 | 0.93 | 0.76 | 0.94251 | 0.67 | 0.7 | 0.73 | 0.7 | 0.7
| Mean |  | 0.89667 | 0.765 | 0.80833 | 0.72667 | 0.715 | 0.735 | 0.735 | 0.77167 | 0.90333 | 0.77333 | 0.95285 | 0.68833 | 0.7 | 0.745 | 0.7 | 0.75
| Median |  | 0.89 | 0.765 | 0.805 | 0.72 | 0.715 | 0.74 | 0.735 | 0.775 | 0.895 | 0.77 | 0.9537 | 0.685 | 0.7 | 0.745 | 0.705 | 0.75
| Std Dev. |  | 0.015055 | 0.010488 | 0.04535 | 0.010328 | 0.015166 | 0.012247 | 0.020736 | 0.023166 | 0.017512 | 0.015055 | 0.018619 | 0.01472 | 0.0063246 | 0.010488 | 0.026077 | 0.054772
| Rel.Std.Dev. |  | 1.679% | 1.371% | 5.61% | 1.421% | 2.121% | 1.666% | 2.821% | 3.002% | 1.939% | 1.947% | 1.954% | 2.138% | 0.9035% | 1.408% | 3.725% | 7.303%
| PDM3 |  | 21.62% | 3.758% | 9.635% | -1.441% | -3.023% | -0.3108% | -0.3108% | 4.662% | 22.52% | 4.888% | 29.24% | -6.64% | -5.058% | 1.045% | -5.058% | 1.724%
| Z-Score (Absolute) |  | 3.41 | 0.45 | 1.42 | 0.41 | 0.67 | 0.22 | 0.22 | 0.6 | 3.56 | 0.64 | 4.67 | 1.27 | 1.01 | 0 | 1.01 | NA
Results from laboratory 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 610 (Certified Value 0.972 wt.%) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*AAS | 3
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 3
| 1 | 1 | 1.05 | > 1 | 0.9521 | 0.971 | 0.956 | 0.968 | 0.975 | 0.949 |  | 0.998 | 1.001 | 0.97791 | 0.9245 | 0.99409 | 0.988 | 1.0047 | 0.9703 | 1.04 | 0.96954 | 0.9346 | 0.948 | 0.93697 | 0.8656 | 1.011 | 1
| 1 | 2 | 1.02 | 0.9821 | 0.94775 | 0.979 | 0.968 | 0.968 | 0.97 | 0.964 |  | 1.01 | 0.99087 | 0.9787 | 0.9207 | 0.96829 | 0.992 | 1.006 | 0.9721 | 1.04 | 0.96311 | 0.9609 | 0.929 | 0.93199 | 0.8613 | 1.038 | 23
| 1 | 3 | 1.02 | 0.9927 | 0.94497 | 0.972 | 0.954 | 0.976 | 0.961 | 0.982 |  | 1.02 | 0.96602 | 0.97386 | 0.91 | 0.98359 | 0.988 | 1.0097 | 0.9727 | 1.06 | 0.9613 | 0.9286 | 0.959 | 0.94473 | 0.8609 | 1.065 | 16
| 1 | 4 | 1.04 | 0.9785 | 0.93798 | 0.959 | 0.942 | 0.966 | 0.971 | 0.982 |  | 0.996 | 0.97363 | 0.98517 | 0.9403 | 0.97852 | 0.993 | 0.99795 | 0.9715 | 1.04 | 0.96946 | 0.9555 | 0.937 | 0.94848 | 0.857 | 1.092 | 0.97152
| 1 | 5 | 1.01 | > 1 | 0.93897 | 0.988 | 0.973 | 0.972 | 0.958 |  | 0.975 | 1.01 | 0.96871 | 0.96016 | 0.9165 | 0.99286 | 1.02 | 0.99552 | 0.9695 | 1.06 | 0.96459 | 0.8799 | 0.947 | 0.94281 | 0.8478 | 1.019 | 80
| 1 | 6 | 1.05 | > 1 | 0.96044 | 0.96 | 0.928 | 0.958 | 0.98 |  |  | 1.03 | 0.97032 | 0.96569 | 0.9464 | 0.96917 | 0.975 | 1.0057 | 0.9703 | 1.07 | 0.96001 | 0.9606 | 0.95 | 0.95342 | 0.8601 | 1.046
| Mean |  | 1.0317 | 0.98443 | 0.94704 | 0.9715 | 0.9535 | 0.968 | 0.96917 | 0.96925 | 0.975 | 1.0107 | 0.97842 | 0.97358 | 0.9264 | 0.98109 | 0.99267 | 1.0033 | 0.97107 | 1.0517 | 0.96467 | 0.93668 | 0.945 | 0.94307 | 0.85878 | 1.0452
| Median |  | 1.03 | 0.9821 | 0.94636 | 0.9715 | 0.955 | 0.968 | 0.9705 | 0.973 | 0.975 | 1.01 | 0.97197 | 0.97589 | 0.9226 | 0.98105 | 0.99 | 1.0052 | 0.9709 | 1.05 | 0.96385 | 0.94505 | 0.9475 | 0.94377 | 0.8605 | 1.042
| Std Dev. |  | 0.017224 | 0.007382 | 0.0084535 | 0.011113 | 0.01661 | 0.0060663 | 0.0083287 | 0.015945 |  | 0.012941 | 0.014149 | 0.0091845 | 0.014116 | 0.01119 | 0.014855 | 0.0053846 | 0.0012291 | 0.013292 | 0.004054 | 0.031005 | 0.010526 | 0.0077355 | 0.006047 | 0.029969
| Rel.Std.Dev. |  | 1.67% | 0.7499% | 0.8926% | 1.144% | 1.742% | 0.6267% | 0.8594% | 1.645% |  | 1.28% | 1.446% | 0.9434% | 1.524% | 1.141% | 1.496% | 0.5367% | 0.1266% | 1.264% | 0.4202% | 3.31% | 1.114% | 0.8202% | 0.7041% | 2.867%
| PDM3 |  | 6.191% | 1.329% | -2.521% | -0.002478% | -1.855% | -0.3627% | -0.2427% | -0.2341% | 0.3578% | 4.029% | 0.71% | 0.2118% | -4.645% | 0.9842% | 2.176% | 3.266% | -0.04708% | 8.249% | -0.7057% | -3.586% | -2.73% | -2.929% | -11.6% | 7.58%
| Z-Score (Absolute) |  | 2.08 | 0.45 | 0.83 | 0.01 | 0.61 | 0.11 | 0.07 | 0.07 | 0.13 | 1.36 | 0.25 | 0.08 | 1.55 | 0.34 | 0.74 | 1.1 | 0.01 | 2.77 | 0.23 | 1.19 | 0.91 | 0.97 | 3.87 | 2.54
| Analytical results for Dy in OREAS 610 (Indicative Value 0.82 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 0.9 | 1
| 1 | 2 | 2.01: 0.9 | 10
| 1 | 3 | 2.01: 0.8 | 16
| 1 | 4 | 2.01: 0.8 | 0.81667
| 1 | 5 | 2.01: 0.7 | 16
| 1 | 6 | 2.01: 0.8
| Mean |  | 2.01: 0.81667
| Median |  | 2.01: 0.8
| Std Dev. |  | 2.01: 0.075277
| Rel.Std.Dev. |  | 2.01: 9.218%
| PDM3 |  | 2.01: -2.22e-14%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Er in OREAS 610 (Indicative Value 0.27 ppm) | [25 empty] | Indicative
Er | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 0.3 | 1
| 1 | 2 | 2.01: 0.3 | 11
| 1 | 3 | 2.01: 0.2 | 16
| 1 | 4 | 2.01: 0.3 | 0.26667
| 1 | 5 | 2.01: 0.2 | 17
| 1 | 6 | 2.01: 0.3
| Mean |  | 2.01: 0.26667
| Median |  | 2.01: 0.3
| Std Dev. |  | 2.01: 0.05164
| Rel.Std.Dev. |  | 2.01: 19.36%
| PDM3 |  | 2.01: -1.221e-13%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Eu in OREAS 610 (Indicative Value 0.23 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 0.3 | 1
| 1 | 2 | 2.01: 0.2 | 12
| 1 | 3 | 2.01: 0.2 | 16
| 1 | 4 | 2.01: 0.2 | 0.23333
| 1 | 5 | 2.01: 0.2 | 18
| 1 | 6 | 2.01: 0.3
| Mean |  | 2.01: 0.23333
| Median |  | 2.01: 0.2
| Std Dev. |  | 2.01: 0.05164
| Rel.Std.Dev. |  | 2.01: 22.13%
| PDM3 |  | 2.01: 1.332e-13%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Fe in OREAS 610 (Certified Value 2.27 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 2
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 3
| 1 | 1 | 2.73 | 2.51 | 2.25 | 2.35 | 2.15 | 2.3 | 2.22 | 2.18 | 2.2 | 2.47 | 2.295 | 2.3609 | 2.16 | 2.24 | 2.34 | 2.1986 | 2.31 | 2.298 | 2.27 | 2.38 | 2.39 | 2 | 2.26 | 2.477 | 1
| 1 | 2 | 2.65 | 2.39 | 2.28 | 2.36 | 2.18 | 2.32 | 2.21 | 2.18 | 2.1 | 2.5 | 2.247 | 2.455 | 2.14 | 2.19 | 2.36 | 2.2014 | 2.32 | 2.337 | 2.27 | 2.43 | 2.38 | 2 | 2.25 | 2.441 | #N/A
| 1 | 3 | 2.63 | 2.45 | 2.28 | 2.25 | 2.17 | 2.28 | 2.18 | 2.22 | 2.17 | 2.51 | 2.111 | 2.4111 | 2.12 | 2.21 | 2.24 | 2.2009 | 2.32 | 2.324 | 2.29 | 2.39 | 2.41 | 2.01 | 2.24 | 2.46 | 16
| 1 | 4 | 2.74 | 2.4 | 2.28 | 2.34 | 2.13 | 2.3 | 2.2 | 2.17 | 2.17 | 2.5 | 2.105 | 2.3637 | 2.18 | 2.23 | 2.29 | 2.184 | 2.32 | 2.375 | 2.24 | 2.38 | 2.39 | 2 | 2.2 | 2.469 | 2.2728
| 1 | 5 | 2.65 | 2.47 | 2.28 | 2.14 | 2.16 | 2.29 | 2.18 | 2.11 | 2.28 | 2.48 | 2.107 | 2.4034 | 2.16 | 2.23 | 2.18 | 2.1902 | 2.32 | 2.364 | 2.24 | 2.21 | 2.39 | 1.99 | 2.19 | 2.433 | 81
| 1 | 6 | 2.69 | 2.47 | 2.26 | 2.26 | 2.13 | 2.28 | 2.2 | 2.12 | 2.2 | 2.54 | 2.187 | 2.4208 | 2.18 | 2.2 | 2.14 | 2.1944 | 2.29 | 2.354 | 2.28 | 2.41 | 2.33 | 2.02 | 2.18 | 2.397
| Mean |  | 2.6817 | 2.4483 | 2.2717 | 2.2833 | 2.1533 | 2.295 | 2.1983 | 2.1633 | 2.1867 | 2.5 | 2.1753 | 2.4025 | 2.1567 | 2.2167 | 2.2583 | 2.1949 | 2.3133 | 2.342 | 2.265 | 2.3667 | 2.3817 | 2.0033 | 2.22 | 2.4462
| Median |  | 2.67 | 2.46 | 2.28 | 2.3 | 2.155 | 2.295 | 2.2 | 2.175 | 2.185 | 2.5 | 2.149 | 2.4073 | 2.16 | 2.22 | 2.265 | 2.1965 | 2.32 | 2.3455 | 2.27 | 2.385 | 2.39 | 2 | 2.22 | 2.4505
| Std Dev. |  | 0.045789 | 0.045789 | 0.013292 | 0.084538 | 0.020656 | 0.015166 | 0.016021 | 0.041312 | 0.058538 | 0.024495 | 0.081667 | 0.035773 | 0.023381 | 0.019664 | 0.087731 | 0.0068196 | 0.012111 | 0.028291 | 0.020736 | 0.079162 | 0.027142 | 0.010328 | 0.034059 | 0.02926
| Rel.Std.Dev. |  | 1.707% | 1.87% | 0.5851% | 3.702% | 0.9593% | 0.6608% | 0.7288% | 1.91% | 2.677% | 0.9798% | 3.754% | 1.489% | 1.084% | 0.8871% | 3.885% | 0.3107% | 0.5235% | 1.208% | 0.9155% | 3.345% | 1.14% | 0.5155% | 1.534% | 1.196%
| PDM3 |  | 17.99% | 7.724% | -0.04898% | 0.4643% | -5.256% | 0.9777% | -3.276% | -4.816% | -3.789% | 9.997% | -4.288% | 5.707% | -5.109% | -2.469% | -0.6356% | -3.425% | 1.784% | 3.046% | -0.3423% | 4.131% | 4.791% | -11.86% | -2.322% | 7.629%
| Z-Score (Absolute) |  | 3.19 | 1.39 | 0.03 | 0.12 | 0.89 | 0.21 | 0.54 | 0.81 | 0.63 | 1.79 | 0.72 | 1.04 | 0.86 | 0.4 | 0.08 | 0.57 | 0.35 | 0.57 | 0.03 | 0.76 | 0.88 | 2.05 | 0.37 | 1.37
| Analytical results for Ga in OREAS 610 (Certified Value 6.36 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 01g |  |  |  | 3
| 1 | 1 | 6.9 | 6.5 | 8.41 | 6.22 | 6.35 | 6.27 | 6.16 | 6.06 | 6 | 6.4 | 7.652 | 7 | 7.1 | 6.37 | 6.5239 | 5 | 6 | 6.7 | 6 | 7 | 8.639 |  |  |  | 1
| 1 | 2 | 6.8 | 6.4 | 8.51 | 5.76 | 6.41 | 6.55 | 6.11 | 6.08 | 6.1 | 6.2 | 7.6627 | 8 | 7.3 | 6.4 | 6.411 | 5.1 | 6 | 6.6 | 6 | 7 | 6.86 |  |  |  | 41
| 1 | 3 | 6.5 | 6 | 7.92 | 6 | 6.35 | 6.43 | 6 | 6.09 | 6 | 6.1 | 7.4639 | 9 | 7.4 | 6.22 | 6.0761 | 5.04 | 6.1 | 6.7 | 6 | 7 | 5.081 |  |  |  | 16
| 1 | 4 | 7 | 6.3 | 8.13 | 5.92 | 6.29 | 6.69 | 6.02 | 6.16 | 6.1 | 6.1 | 7.5565 | 8 | 7.1 | 6.33 | 5.8041 | 4.87 | 6.1 | 6.6 | 6 | 6 | 7.6 |  |  |  | 6.3593
| 1 | 5 | 6.8 | 6.5 | 7.76 | 5.82 | 6.23 | 6.43 | 5.63 | 6.33 | 6.3 | 6.1 | 7.5088 | 8 | 7.7 | 6.54 | 6.3434 | 5.07 | 6 | 6 | 6 | 7 | 6.9 |  |  |  | 82
| 1 | 6 | 6.9 | 6.6 | 8.03 | 5.78 | 6.28 | 6.51 | 6.03 | 6.21 | 6.2 | 6.5 | 7.4723 | 8 | 6.7 | 6.17 | 5.5114 | 5.1 | 6.1 | 6.5 | 6 | 7 | 6.4
| Mean |  | 6.8167 | 6.3833 | 8.1267 | 5.9167 | 6.3183 | 6.48 | 5.9917 | 6.155 | 6.1167 | 6.2333 | 7.5527 | 8 | 7.2167 | 6.3383 | 6.1117 | 5.03 | 6.05 | 6.5167 | 6 | 6.8333 | 6.9133
| Median |  | 6.85 | 6.45 | 8.08 | 5.87 | 6.32 | 6.47 | 6.025 | 6.125 | 6.1 | 6.15 | 7.5327 | 8 | 7.2 | 6.35 | 6.2098 | 5.055 | 6.05 | 6.6 | 6 | 7 | 6.88
| Std Dev. |  | 0.17224 | 0.2137 | 0.28766 | 0.17409 | 0.064005 | 0.14071 | 0.18734 | 0.10252 | 0.1169 | 0.17512 | 0.087477 | 0.63246 | 0.33714 | 0.13258 | 0.39228 | 0.087178 | 0.054772 | 0.26394 | 0 | 0.40825 | 1.1894
| Rel.Std.Dev. |  | 2.527% | 3.348% | 3.54% | 2.942% | 1.013% | 2.171% | 3.127% | 1.666% | 1.911% | 2.809% | 1.158% | 7.906% | 4.672% | 2.092% | 6.419% | 1.733% | 0.9053% | 4.05% | 0% | 5.974% | 17.2%
| PDM3 |  | 7.193% | 0.3786% | 27.79% | -6.96% | -0.6435% | 1.899% | -5.78% | -3.212% | -3.815% | -1.98% | 18.77% | 25.8% | 13.48% | -0.329% | -3.894% | -20.9% | -4.863% | 2.475% | -5.649% | 7.455% | 8.713%
| Z-Score (Absolute) |  | 1.33 | 0.15 | 4.9 | 1.12 | 0.03 | 0.41 | 0.92 | 0.47 | 0.58 | 0.26 | 3.34 | NA | 2.42 | 0.03 | 0.59 | 3.54 | 0.76 | 0.51 | NA | NA | 1.59
Results from laboratories 13, 23 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 610 (Indicative Value 1.43 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | 2.02: Lab |  |  |  | 1
| No | No | 2.02: 2.04 |  |  |  | ppm
|  |  | 2.02: AR*MS |  |  |  | 2
|  |  | 2.02: 0.5g |  |  |  | 2
| 1 | 1 | 2.02: 1.6 |  |  |  | 1
| 1 | 2 | 2.02: 1.5 |  |  |  | 13
| 1 | 3 | 2.02: 1.4 |  |  |  | 16
| 1 | 4 | 2.02: 1.3 |  |  |  | 1.4333
| 1 | 5 | 2.02: 1.4 |  |  |  | 19
| 1 | 6 | 2.02: 1.4
| Mean |  | 2.02: 1.4333
| Median |  | 2.02: 1.4
| Std Dev. |  | 2.02: 0.10328
| Rel.Std.Dev. |  | 2.02: 7.206%
| PDM3 |  | 2.02: 2.22e-13%
Z-Score (Absolute) | bar: 2.02
| Analytical results for Ge in OREAS 610 (Indicative Value 0.28 ppm) | [25 empty] | Indicative
Ge | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.08; 2.08: 2.09; 2.09: 2.11; 2.10: 2.18; 2.11: 2.20; 2.12: 2.21; 2.13: 2.23 |  |  |  | ppm
|  |  | 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*OES/MS; 2.11: AR*OES/MS; 2.12: AR*OES/MS; 2.13: AR*OES/MS |  |  |  | 2
|  |  | 2.02: 15g; 2.03: 0.5g; 2.04: 0.5g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 0.5g; 2.10: 0.2g; 2.11: 0.25g; 2.12: 0.25g; 2.13: 0.25g |  |  |  | 2
| 1 | 1 | 2.02: 0.7; 2.03: 0.4; 2.04: 0.45; 2.05: 0.24; 2.06: 0.32; 2.07: 0.2; 2.08: 0.18; 2.09: 0.4; 2.10: 0.11; 2.11: 0.6; 2.12: 0.7; 2.13: < 0.1 |  |  |  | 1
| 1 | 2 | 2.02: 0.6; 2.03: 0.4; 2.04: 0.46; 2.05: 0.26; 2.06: 0.4; 2.07: 0.19; 2.08: 0.16; 2.09: 0.35; 2.10: 0.12; 2.11: 0.7; 2.12: 0.7; 2.13: < 0.1 |  |  |  | 6
| 1 | 3 | 2.02: 0.7; 2.03: 0.3; 2.04: 0.44; 2.05: 0.24; 2.06: 0.37; 2.07: 0.22; 2.08: 0.16; 2.09: 0.45; 2.10: 0.11; 2.11: 0.7; 2.12: 0.6; 2.13: < 0.1 |  |  |  | 16
| 1 | 4 | 2.02: 0.8; 2.03: 0.3; 2.04: 0.41; 2.05: 0.28; 2.06: 0.37; 2.07: 0.21; 2.08: 0.17; 2.09: 0.41; 2.10: 0.11; 2.11: 0.7; 2.12: 0.7; 2.13: < 0.1 |  |  |  | 0.28419
| 1 | 5 | 2.02: 0.6; 2.03: 0.5; 2.04: 0.47; 2.05: 0.31; 2.06: 0.39; 2.07: 0.26; 2.08: 0.14; 2.09: 0.4; 2.10: 0.11; 2.11: 0.7; 2.12: 0.6; 2.13: < 0.1 |  |  |  | 20
| 1 | 6 | 2.02: 0.8; 2.03: 0.6; 2.04: 0.44; 2.05: 0.28; 2.06: 0.43; 2.07: 0.48; 2.08: 0.17; 2.09: 0.42; 2.10: 0.11; 2.11: 0.7; 2.12: 0.6; 2.13: < 0.1
| Mean |  | 2.02: 0.7; 2.03: 0.41667; 2.04: 0.445; 2.05: 0.26833; 2.06: 0.38; 2.07: 0.26; 2.08: 0.16333; 2.09: 0.405; 2.10: 0.11167; 2.11: 0.68333; 2.12: 0.65
| Median |  | 2.02: 0.7; 2.03: 0.4; 2.04: 0.445; 2.05: 0.27; 2.06: 0.38; 2.07: 0.215; 2.08: 0.165; 2.09: 0.405; 2.10: 0.11; 2.11: 0.7; 2.12: 0.65
| Std Dev. |  | 2.02: 0.089443; 2.03: 0.1169; 2.04: 0.020736; 2.05: 0.027142; 2.06: 0.036878; 2.07: 0.11045; 2.08: 0.013663; 2.09: 0.032711; 2.10: 0.0040825; 2.11: 0.040825; 2.12: 0.054772
| Rel.Std.Dev. |  | 2.02: 12.78%; 2.03: 28.06%; 2.04: 4.66%; 2.05: 10.11%; 2.06: 9.705%; 2.07: 42.48%; 2.08: 8.365%; 2.09: 8.077%; 2.10: 3.656%; 2.11: 5.974%; 2.12: 8.427%
| PDM3 |  | 2.02: 146.3%; 2.03: 46.62%; 2.04: 56.59%; 2.05: -5.58%; 2.06: 33.71%; 2.07: -8.512%; 2.08: -42.53%; 2.09: 42.51%; 2.10: -60.71%; 2.11: 140.4%; 2.12: 128.7%
| Z-Score (Absolute) |  | 2.02: NA; 2.03: NA; 2.04: 0.95; 2.05: 0.02; 2.06: 0.61; 2.07: 0.02; 2.08: 0.53; 2.09: 0.74; 2.10: 0.8; 2.11: NA; 2.12: NA; 2.13: 1.13
Results from laboratories 3, 4, 20 & 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 610 (Certified Value 0.38 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 0.37 | 0.4 | 0.5 | 0.39 | 0.41 | 0.41 | 0.44 | 0.42 | 0.36 | 0.4 | 0.34376 | 0.35 | 0.35 | 0.4 | 0.36 | 0.5 | [8 empty] | 1
| 1 | 2 | 0.36 | 0.36 | 0.5 | 0.42 | 0.43 | 0.41 | 0.42 | 0.4 | 0.36 | 0.39 | 0.36493 | 0.34 | 0.37 | 0.38 | 0.39 | 0.5 | [8 empty] | 27
| 1 | 3 | 0.36 | 0.38 | 0.4 | 0.39 | 0.41 | 0.41 | 0.4 | 0.43 | 0.36 | 0.4 | 0.35449 | 0.36 | 0.36 | 0.37 | 0.35 | 0.5 | [8 empty] | 16
| 1 | 4 | 0.37 | 0.42 | 0.5 | 0.37 | 0.43 | 0.42 | 0.41 | 0.43 | 0.35 | 0.39 | 0.35423 | 0.35 | 0.34 | 0.39 | 0.34 | 0.5 | [8 empty] | 0.38357
| 1 | 5 | 0.37 | 0.38 | 0.4 | 0.4 | 0.41 | 0.41 | 0.41 | 0.42 | 0.37 | 0.39 | 0.34555 | 0.34 | 0.36 | 0.4 | 0.33 | 0.5 | [8 empty] | 83
| 1 | 6 | 0.38 | 0.41 | 0.5 | 0.4 | 0.41 | 0.4 | 0.4 | 0.45 | 0.35 | 0.39 | 0.35675 | 0.35 | 0.35 | 0.38 | 0.36 | 0.5
| Mean |  | 0.36833 | 0.39167 | 0.46667 | 0.395 | 0.41667 | 0.41 | 0.41333 | 0.425 | 0.35833 | 0.39333 | 0.35328 | 0.34833 | 0.355 | 0.38667 | 0.355 | 0.5
| Median |  | 0.37 | 0.39 | 0.5 | 0.395 | 0.41 | 0.41 | 0.41 | 0.425 | 0.36 | 0.39 | 0.35436 | 0.35 | 0.355 | 0.385 | 0.355 | 0.5
| Std Dev. |  | 0.0075277 | 0.022286 | 0.05164 | 0.016432 | 0.010328 | 0.0063246 | 0.015055 | 0.016432 | 0.0075277 | 0.005164 | 0.0077541 | 0.0075277 | 0.010488 | 0.012111 | 0.020736 | 0
| Rel.Std.Dev. |  | 2.044% | 5.69% | 11.07% | 4.16% | 2.479% | 1.543% | 3.642% | 3.866% | 2.101% | 1.313% | 2.195% | 2.161% | 2.954% | 3.132% | 5.841% | 0%
| PDM3 |  | -3.972% | 2.111% | 21.66% | 2.98% | 8.629% | 6.891% | 7.76% | 10.8% | -6.579% | 2.546% | -7.895% | -9.186% | -7.448% | 0.8079% | -7.448% | 30.36%
| Z-Score (Absolute) |  | 0.54 | 0.07 | NA | 0.15 | 0.72 | 0.54 | 0.63 | 0.94 | 0.8 | 0.11 | 0.94 | 1.07 | 0.89 | 0.07 | 0.89 | NA
Results from laboratories 4 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 610 (Certified Value 0.8 ppm) | [25 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 10g | 0.25g | 01g | [6 empty] | 3
| 1 | 1 | 0.733 | 0.64 | 0.79 | 0.78 | 0.81 | 0.76 | 0.79 | 0.82 | 0.60254 | 1 | 0.929 | < 2 | < 1 | 0.86 | 0.8 | < 5 | 0.852 | < 5 | [6 empty] | 1
| 1 | 2 | 0.755 | 0.62 | 0.82 | 0.81 | 0.79 | 0.75 | 0.74 | 0.81 | 0.73248 | < 1 | 0.897 | < 2 | < 1 | 0.86 | 0.84 | < 5 | 0.86 | < 5 | [6 empty] | 5
| 1 | 3 | 0.766 | 0.61 | 0.74 | 0.81 | 0.83 | 0.76 | 0.81 | 0.83 | 0.69772 | 1 | 0.867 | < 2 | < 1 | 0.81 | 0.79 | < 5 | 0.896 | < 5 | [6 empty] | 16
| 1 | 4 | 0.777 | 0.67 | 0.76 | 0.79 | 0.86 | 0.77 | 0.77 | 0.82 | 0.65534 | < 1 | 0.904 | < 2 | < 1 | 0.82 | 0.74 | < 5 | 0.869 | < 5 | [6 empty] | 0.80046
| 1 | 5 | 0.759 | 0.57 | 0.78 | 0.81 | 0.84 | 0.76 | 0.84 | 0.83 | 0.65393 | < 1 | 0.899 | < 2 | < 1 | 0.87 | 0.73 | < 5 | 0.855 | < 5 | [6 empty] | 84
| 1 | 6 | 0.773 | 0.6 | 0.75 | 0.81 | 0.84 | 0.76 | 0.77 | 0.81 | 0.65617 | 1 | 0.923 | < 2 | < 1 | 0.86 | 0.79 | < 5 | 0.828 | < 5
| Mean |  | 0.7605 | 0.61833 | 0.77333 | 0.80167 | 0.82833 | 0.76 | 0.78667 | 0.82 | 0.66636 | 1 | 0.90317 |  |  | 0.84667 | 0.78167 |  | 0.86
| Median |  | 0.7625 | 0.615 | 0.77 | 0.81 | 0.835 | 0.76 | 0.78 | 0.82 | 0.65575 | 1 | 0.9015 |  |  | 0.86 | 0.79 |  | 0.8575
| Std Dev. |  | 0.015796 | 0.034303 | 0.029439 | 0.013292 | 0.024833 | 0.0063246 | 0.035024 | 0.0089443 | 0.044289 | 0 | 0.022022 |  |  | 0.025033 | 0.040702 |  | 0.022316
| Rel.Std.Dev. |  | 2.077% | 5.548% | 3.807% | 1.658% | 2.998% | 0.8322% | 4.452% | 1.091% | 6.646% | 0% | 2.438% |  |  | 2.957% | 5.207% |  | 2.595%
| PDM3 |  | -4.992% | -22.75% | -3.388% | 0.1514% | 3.483% | -5.054% | -1.723% | 2.442% | -16.75% | 24.93% | 12.83% |  |  | 5.773% | -2.347% |  | 7.439%
| Z-Score (Absolute) |  | 0.62 | 2.75 | 0.42 | 0 | 0.4 | 0.62 | 0.22 | 0.27 | 2.03 | NA | 1.52 | 2.97 | 4.52 | 0.67 | 0.3 | 25.45 | 0.87 | 25.45
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 610 (Indicative Value 0.092 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | 2.03: Lab | [6 empty] | 1
| No | No | 2.03: 2.04 | [6 empty] | ppm
|  |  | 2.03: AR*MS | [6 empty] | 3
|  |  | 2.03: 0.5g | [6 empty] | 3
| 1 | 1 | 2.03: < 0.1 | [6 empty] | 1
| 1 | 2 | 2.03: 0.1 | [6 empty] | 8
| 1 | 3 | 2.03: 0.1 | [6 empty] | 16
| 1 | 4 | 2.03: 0.1 | [6 empty] | 0.091667
| 1 | 5 | 2.03: 0.1 | [6 empty] | 14
| 1 | 6 | 2.03: 0.1
| Mean |  | 2.03: 0.1
| Median |  | 2.03: 0.1
| Std Dev. |  | 2.03: 0
Rel.Std.Dev. | bar: 2.03
| PDM3 |  | 2.03: 9.091%
Z-Score (Absolute) | bar: 2.03
| Analytical results for In in OREAS 610 (Certified Value 3.76 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 3.3 | 4.1 | 3.85 | 3.81 | 3.94 | 3.66 | 4.26 | 3.74 | 3.74 | 3.8 | 3.8387 | 3.9 | 5 | 3.85 | 3.52 | 4.1 | [8 empty] | 1
| 1 | 2 | 3.3 | 4.12 | 3.89 | 3.71 | 3.96 | 3.78 | 3.95 | 3.61 | 3.8 | 3.74 | 3.8399 | 3.78 | 5.1 | 3.81 | 3.53 | 4 | [8 empty] | 29
| 1 | 3 | 3.27 | 4.1 | 3.7 | 3.67 | 3.93 | 3.72 | 4.08 | 3.71 | 3.94 | 3.75 | 3.5675 | 3.73 | 5.1 | 3.77 | 3.64 | 4.1 | [8 empty] | 16
| 1 | 4 | 3.29 | 3.98 | 3.72 | 3.68 | 3.84 | 3.88 | 3.95 | 3.75 | 3.7 | 3.8 | 3.9393 | 3.87 | 5 | 3.71 | 3.6 | 4.1 | [8 empty] | 3.7628
| 1 | 5 | 3.36 | 4.03 | 3.57 | 3.74 | 3.81 | 3.69 | 3.78 | 3.92 | 3.82 | 3.75 | 3.6574 | 3.92 | 5 | 3.7 | 3.3 | 4 | [8 empty] | 85
| 1 | 6 | 3.37 | 3.81 | 3.87 | 3.73 | 3.92 | 3.73 | 4.03 | 3.88 | 3.76 | 3.72 | 3.7239 | 3.72 | 5.1 | 3.82 | 3.6 | 4.1
| Mean |  | 3.315 | 4.0233 | 3.7667 | 3.7233 | 3.9 | 3.7433 | 4.0083 | 3.7683 | 3.7933 | 3.76 | 3.7611 | 3.82 | 5.05 | 3.7767 | 3.5317 | 4.0667
| Median |  | 3.3 | 4.065 | 3.785 | 3.72 | 3.925 | 3.725 | 3.99 | 3.745 | 3.78 | 3.75 | 3.7813 | 3.825 | 5.05 | 3.79 | 3.565 | 4.1
| Std Dev. |  | 0.040373 | 0.11708 | 0.12501 | 0.050465 | 0.060332 | 0.078145 | 0.15993 | 0.11409 | 0.083586 | 0.032863 | 0.13686 | 0.087864 | 0.054772 | 0.06121 | 0.12238 | 0.05164
| Rel.Std.Dev. |  | 1.218% | 2.91% | 3.319% | 1.355% | 1.547% | 2.088% | 3.99% | 3.028% | 2.204% | 0.874% | 3.639% | 2.3% | 1.085% | 1.621% | 3.465% | 1.27%
| PDM3 |  | -11.9% | 6.924% | 0.1028% | -1.049% | 3.646% | -0.5173% | 6.525% | 0.1471% | 0.8115% | -0.07436% | -0.04483% | 1.52% | 34.21% | 0.3686% | -6.143% | 8.076%
| Z-Score (Absolute) |  | 6.38 | 3.5 | 0.08 | 0.69 | 1.78 | 0.41 | 3.29 | 0.06 | 0.29 | 0.17 | 0.16 | 0.66 | 17.82 | 0.06 | 3.36 | 4.1
| Analytical results for K in OREAS 610 (Certified Value 0.213 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.329 | 0.22 | 0.32 | 0.2 | 0.23 | 0.2 | 0.22 | 0.22 | 0.288 | 0.227 | 0.31988 | 0.21 | 0.23 | 0.19 | 0.23776 | 0.19 | 0.225 | 0.23 | 0.22 | 0.21 | 0.2 | 0.23 | 0.2681 |  | 1
| 1 | 2 | 0.3246 | 0.23 | 0.33 | 0.21 | 0.23 | 0.21 | 0.22 | 0.22 | 0.3 | 0.224 | 0.31812 | 0.18 | 0.21 | 0.19 | 0.23964 | 0.19 | 0.226 | 0.21 | 0.23 | 0.21 | 0.2 | 0.23 | 0.2375 |  | #N/A
| 1 | 3 | 0.3203 | 0.23 | 0.3 | 0.21 | 0.22 | 0.2 | 0.22 | 0.22 | 0.303 | 0.192 | 0.31718 | 0.18 | 0.23 | 0.17 | 0.24991 | 0.19 | 0.225 | 0.22 | 0.22 | 0.22 | 0.2 | 0.23 | 0.2569 |  | 16
| 1 | 4 | 0.3238 | 0.22 | 0.31 | 0.21 | 0.22 | 0.2 | 0.21 | 0.22 | 0.289 | 0.192 | 0.31029 | 0.19 | 0.23 | 0.18 | 0.22098 | 0.19 | 0.227 | 0.21 | 0.22 | 0.21 | 0.19 | 0.22 | 0.2763 |  | 0.21343
| 1 | 5 | 0.3264 | 0.23 | 0.3 | 0.2 | 0.22 | 0.2 | 0.21 | 0.22 | 0.292 | 0.19 | 0.30742 | 0.19 | 0.23 | 0.17 | 0.21617 | 0.19 | 0.222 | 0.22 | 0.2 | 0.22 | 0.18 | 0.22 | 0.2457 |  | 86
| 1 | 6 | 0.3247 | 0.23 | 0.31 | 0.2 | 0.22 | 0.21 | 0.21 | 0.22 | 0.293 | 0.214 | 0.31571 | 0.19 | 0.23 | 0.17 | 0.21117 | 0.19 | 0.222 | 0.22 | 0.22 | 0.21 | 0.18 | 0.22 | 0.24151
| Mean |  | 0.3248 | 0.22667 | 0.31167 | 0.205 | 0.22333 | 0.20333 | 0.215 | 0.22 | 0.29417 | 0.2065 | 0.31477 | 0.19 | 0.22667 | 0.17833 | 0.22927 | 0.19 | 0.2245 | 0.21833 | 0.21833 | 0.21333 | 0.19167 | 0.225 | 0.25433
| Median |  | 0.32465 | 0.23 | 0.31 | 0.205 | 0.22 | 0.2 | 0.215 | 0.22 | 0.2925 | 0.203 | 0.31645 | 0.19 | 0.23 | 0.175 | 0.22937 | 0.19 | 0.225 | 0.22 | 0.22 | 0.21 | 0.195 | 0.225 | 0.2513
| Std Dev. |  | 0.002881 | 0.005164 | 0.01169 | 0.0054772 | 0.005164 | 0.005164 | 0.0054772 | 0 | 0.006047 | 0.017178 | 0.0048596 | 0.010954 | 0.008165 | 0.0098319 | 0.015321 | 3.0405e-17 | 0.0020736 | 0.0075277 | 0.0098319 | 0.005164 | 0.0098319 | 0.0054772 | 0.015498
| Rel.Std.Dev. |  | 0.887% | 2.278% | 3.751% | 2.672% | 2.312% | 2.54% | 2.548% | 0% | 2.056% | 8.319% | 1.544% | 5.766% | 3.602% | 5.513% | 6.683% | 1.6e-14% | 0.9237% | 3.448% | 4.503% | 2.421% | 5.13% | 2.434% | 6.094%
| PDM3 |  | 52.18% | 6.201% | 46.03% | -3.951% | 4.639% | -4.732% | 0.7345% | 3.077% | 37.83% | -3.248% | 47.48% | -10.98% | 6.201% | -16.45% | 7.421% | -10.98% | 5.186% | 2.296% | 2.296% | -0.04636% | -10.2% | 5.42% | 19.16%
| Z-Score (Absolute) |  | 5.23 | 0.33 | 4.58 | 0.75 | 0.17 | 0.83 | 0.25 | 0 | 3.7 | 0.67 | 4.73 | 1.5 | 0.33 | 2.08 | 0.46 | 1.5 | 0.22 | 0.08 | 0.08 | 0.33 | 1.42 | 0.25 | 1.71
| Analytical results for La in OREAS 610 (Certified Value 6.68 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  | 3
| 1 | 1 | 6.6 | 7.7 | 6.4 | 6.5 | 6.3 | 7 | 6.6 | 11.3 | 6.782 | 9.6259 | 5 | 6.8 | 5.8 | 6.4102 | 8.4 | 7.26 | 7.3 | 6.9 | 6.4 | 8 | 5 | 7.477 |  |  | 1
| 1 | 2 | 6.6 | 7.4 | 6.4 | 6.5 | 6.7 | 6.9 | 6.4 | 11.3 | 6.747 | 9.6628 | 5 | 6.1 | 5.8 | 6.2833 | 8.3 | 7.06 | 7.2 | 7.1 | 6.5 | 7 | 5 | 7.321 |  |  | 30
| 1 | 3 | 6.5 | 7.3 | 6.6 | 6.3 | 6.4 | 6.8 | 6.7 | 11.8 | 7.05 | 9.5635 | 5 | 6.8 | 5.9 | 6.7338 | 8.1 | 7.09 | 7.1 | 6.7 | 6.6 | 7 | 5 | 7.165 |  |  | 16
| 1 | 4 | 6.6 | 7.4 | 6.5 | 6.1 | 6.6 | 6.9 | 6.7 | 10.9 | 6.751 | 9.6585 | 5 | 7.5 | 5.6 | 5.8537 | 8.5 | 7.08 | 7.1 | 6.6 | 6.5 | 7 | 5 | 7.409 |  |  | 6.6766
| 1 | 5 | 6.4 | 6.8 | 6 | 6.4 | 6.6 | 6.5 | 6.9 | 11.4 | 6.865 | 9.5303 | 6 | 6.8 | 5.5 | 5.8892 | 8.4 | 7.09 | 7.3 | 6.2 | 6.4 | 7 | 5 | 7.253 |  |  | 87
| 1 | 6 | 6.5 | 7.4 | 6.1 | 6.3 | 6.8 | 6.8 | 6.9 | 11.7 | 6.655 | 9.6415 | 6 | 6.4 | 5.6 | 5.5143 | 7.9 | 7.07 | 7.2 | 6.6 | 6.4 | 6 | 5 | 7.297
| Mean |  | 6.5333 | 7.3333 | 6.3333 | 6.35 | 6.5667 | 6.8167 | 6.7 | 11.4 | 6.8083 | 9.6137 | 5.3333 | 6.7333 | 5.7 | 6.1141 | 8.2667 | 7.1083 | 7.2 | 6.6833 | 6.4667 | 7 | 5 | 7.3203
| Median |  | 6.55 | 7.4 | 6.4 | 6.35 | 6.6 | 6.85 | 6.7 | 11.35 | 6.7665 | 9.6337 | 5 | 6.8 | 5.7 | 6.0862 | 8.35 | 7.085 | 7.2 | 6.65 | 6.45 | 7 | 5 | 7.309
| Std Dev. |  | 0.08165 | 0.29439 | 0.23381 | 0.15166 | 0.18619 | 0.17224 | 0.18974 | 0.32249 | 0.13628 | 0.054416 | 0.5164 | 0.47188 | 0.15492 | 0.44236 | 0.22509 | 0.075211 | 0.089443 | 0.30605 | 0.08165 | 0.63246 | 0 | 0.11101
| Rel.Std.Dev. |  | 1.25% | 4.014% | 3.692% | 2.388% | 2.835% | 2.527% | 2.832% | 2.829% | 2.002% | 0.566% | 9.682% | 7.008% | 2.718% | 7.235% | 2.723% | 1.058% | 1.242% | 4.579% | 1.263% | 9.035% | 0% | 1.516%
| PDM3 |  | -2.146% | 9.836% | -5.142% | -4.892% | -1.647% | 2.097% | 0.35% | 70.74% | 1.973% | 43.99% | -20.12% | 0.8493% | -14.63% | -8.426% | 23.81% | 6.466% | 7.839% | 0.1004% | -3.145% | 4.843% | -25.11% | 9.641%
| Z-Score (Absolute) |  | 0.35 | 1.06 | 0.7 | 0.67 | 0.29 | 0.15 | 0.06 | 8.21 | 0.13 | 5.07 | NA | 0 | 1.82 | 1.09 | 2.7 | 0.66 | 0.82 | 0.09 | 0.47 | NA | NA | 1.03
Results from laboratories 13, 23 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 610 (Certified Value 8.46 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g |  |  |  |  | 2
| 1 | 1 | 7.3 | 8.9 | 10.8 | 7.9 | 8.4 | 8.1 | 8.2 | 8.6 | 10.7 | 8.32 | 10.1807 | 8.4 | 9.0153 | 7.2 | 6.984 | 9 | 8 | 8 | 8 | 9 |  |  |  |  | 1
| 1 | 2 | 7.1 | 9.3 | 10.5 | 7.9 | 8.4 | 8.3 | 7.8 | 8.4 | 10.7 | 8.19 | 10.3779 | 8 | 9.0281 | 7.7 | 9.101 | 10 | 8 | 8 | 8.2 | 9 |  |  |  |  | #N/A
| 1 | 3 | 6.9 | 8.9 | 10.1 | 8.1 | 8.4 | 8.1 | 7.5 | 8.7 | 10.7 | 8.33 | 10.1288 | 7.7 | 9.1361 | 7.8 | 8.265 | 8 | 8 | 8 | 8.9 | 9 |  |  |  |  | 16
| 1 | 4 | 7.6 | 9.5 | 10.2 | 8.4 | 8.1 | 8.5 | 7.2 | 8.8 | 10.4 | 8.14 | 10.4688 | 8 | 8.7161 | 7.5 | 8.08 | 8 | 8 | 8 | 8.8 | 9 |  |  |  |  | 8.4614
| 1 | 5 | 7.4 | 9.4 | 9.7 | 7.9 | 8.1 | 8.2 | 7 | 9.3 | 10.4 | 8.03 | 10.1783 | 7.6 | 8.7143 | 7.5 | 8.716 | 8 | 7 | 8 | 9 | 9 |  |  |  |  | 88
| 1 | 6 | 7.3 | 8.7 | 10.3 | 7.8 | 8.3 | 8.2 | 8.5 | 9.2 | 10.3 | 8.02 | 10.2892 | 7.4 | 8.6101 | 7.4 | 8.284 | 10 | 8 | 8 | 7.4 | 9
| Mean |  | 7.2667 | 9.1167 | 10.2667 | 8 | 8.2833 | 8.2333 | 7.7 | 8.8333 | 10.5333 | 8.1717 | 10.2706 | 7.85 | 8.87 | 7.5167 | 8.2383 | 8.8333 | 7.8333 | 8 | 8.3833 | 9
| Median |  | 7.3 | 9.1 | 10.25 | 7.9 | 8.35 | 8.2 | 7.65 | 8.75 | 10.55 | 8.165 | 10.2349 | 7.85 | 8.8657 | 7.5 | 8.2745 | 8.5 | 8 | 8 | 8.5 | 9
| Std Dev. |  | 0.24221 | 0.32506 | 0.37238 | 0.21909 | 0.1472 | 0.15055 | 0.57966 | 0.35024 | 0.18619 | 0.13527 | 0.13256 | 0.35637 | 0.21563 | 0.2137 | 0.71744 | 0.98319 | 0.40825 | 0 | 0.62743 | 0
| Rel.Std.Dev. |  | 3.333% | 3.566% | 3.627% | 2.739% | 1.777% | 1.829% | 7.528% | 3.965% | 1.768% | 1.655% | 1.291% | 4.54% | 2.431% | 2.843% | 8.709% | 11.13% | 5.212% | 0% | 7.484% | 0%
| PDM3 |  | -14.12% | 7.745% | 21.34% | -5.453% | -2.104% | -2.695% | -8.998% | 4.396% | 24.49% | -3.424% | 21.38% | -7.225% | 4.829% | -11.16% | -2.636% | 4.396% | -7.422% | -5.453% | -0.9223% | 6.366%
| Z-Score (Absolute) |  | 1.18 | 1.02 | 2.39 | 0.31 | 0.03 | 0.03 | 0.67 | 0.68 | 2.7 | 0.11 | 2.39 | 0.49 | 0.72 | 0.89 | 0.03 | NA | NA | NA | 0.15 | NA
Results from laboratories 20, 21, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 610 (Indicative Value 31.7 ppb) | [25 empty] | Indicative
Lu | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab |  |  |  |  | 1
| No | No | 2.02: 2.04; 2.03: 2.20; 2.04: 2.21 |  |  |  |  | ppb
|  |  | 2.02: AR*MS; 2.03: AR*OES/MS; 2.04: AR*OES/MS |  |  |  |  | 1
|  |  | 2.02: 0.5g; 2.03: 0.25g; 2.04: 0.25g |  |  |  |  | 1
| 1 | 1 | 2.02: < 100; 2.03: 40; 2.04: 30 |  |  |  |  | 1
| 1 | 2 | 2.02: < 100; 2.03: 30; 2.04: 30 |  |  |  |  | 6
| 1 | 3 | 2.02: < 100; 2.03: 40; 2.04: 30 |  |  |  |  | 16
| 1 | 4 | 2.02: < 100; 2.03: 30; 2.04: 30 |  |  |  |  | 31.6667
| 1 | 5 | 2.02: < 100; 2.03: 30; 2.04: 30 |  |  |  |  | 15
| 1 | 6 | 2.02: < 100; 2.03: 40; 2.04: 20
| Mean |  | 2.03: 35; 2.04: 28.3333
| Median |  | 2.03: 35; 2.04: 30
| Std Dev. |  | 2.03: 5.4772; 2.04: 4.0825
| Rel.Std.Dev. |  | 2.03: 15.65%; 2.04: 14.41%
| PDM3 |  | 2.03: 10.53%; 2.04: -10.53%
| Z-Score (Absolute) |  | 2.02: 1.52; 2.03: 0; 2.04: 0.67
| Analytical results for Mg in OREAS 610 (Certified Value 1059 ppm) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 1200 | 1000 | 1200 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 1080 | 1130.44 | 1000 | 1100 | 1100 | 1077.5139 | 1100 | 1050 | 1200 | 1100 | 1000 | 1000 | 1100 | 1223 |  | 1
| 1 | 2 | 1200 | 1000 | 1200 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 1080 | 1204.2877 | 1000 | 1100 | 1100 | 1078.6444 | 1100 | 1040 | 1000 | 1100 | 1000 | 1000 | 1100 | 1211 |  | #N/A
| 1 | 3 | 1200 | 1000 | 1100 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 940 | 1190.6996 | 1000 | 1100 | 1000 | 1081.3329 | 1100 | 1050 | 1100 | 1100 | 1000 | 1000 | 1100 | 1199 |  | 16
| 1 | 4 | 1100 | 1000 | 1100 | 1000 | 1000 | 1000 | 1000 | 1000 | 1200 | 920 | 1112.4571 | 1000 | 1100 | 1100 | 1059.0654 | 1100 | 1040 | 1200 | 1100 | 1000 | 1000 | 1000 | 1187 |  | 1058.8625
| 1 | 5 | 1200 | 1000 | 1100 | 1000 | 1000 | 1000 | 1000 | 1100 | 1300 | 920 | 1173.2879 | 1000 | 1100 | 1000 | 1051.5375 | 1100 | 1040 | 1100 | 1000 | 1000 | 1000 | 1000 | 1175 |  | 89
| 1 | 6 | 1200 | 1000 | 1100 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 1000 | 1201.7996 | 1000 | 1100 | 1000 | 1050.783 | 1100 | 1040 | 1100 | 1100 | 1000 | 1000 | 1000 | 1163
| Mean |  | 1183.3333 | 1000 | 1133.3333 | 1000 | 1000 | 1000 | 1000 | 1016.6667 | 1283.3333 | 990 | 1168.8286 | 1000 | 1100 | 1050 | 1066.4795 | 1100 | 1043.3333 | 1116.6667 | 1083.3333 | 1000 | 1000 | 1050 | 1193
| Median |  | 1200 | 1000 | 1100 | 1000 | 1000 | 1000 | 1000 | 1000 | 1300 | 970 | 1181.9937 | 1000 | 1100 | 1050 | 1068.2896 | 1100 | 1040 | 1100 | 1100 | 1000 | 1000 | 1050 | 1193
| Std Dev. |  | 40.8248 | 0 | 51.6398 | 0 | 0 | 0 | 0 | 40.8248 | 40.8248 | 75.6307 | 38.7146 | 0 | 0 | 54.7723 | 14.2477 | 0 | 5.164 | 75.2773 | 40.8248 | 0 | 0 | 54.7723 | 22.4499
| Rel.Std.Dev. |  | 3.45% | 0% | 4.556% | 0% | 0% | 0% | 0% | 4.016% | 3.181% | 7.639% | 3.312% | 0% | 0% | 5.216% | 1.336% | 0% | 0.4949% | 6.741% | 3.768% | 0% | 0% | 5.216% | 1.882%
| PDM3 |  | 11.76% | -5.559% | 7.033% | -5.559% | -5.559% | -5.559% | -5.559% | -3.985% | 21.2% | -6.503% | 10.39% | -5.559% | 3.885% | -0.837% | 0.7194% | 3.885% | -1.467% | 5.459% | 2.311% | -5.559% | -5.559% | -0.837% | 12.67%
| Z-Score (Absolute) |  | 1.8 | 0.67 | 1.12 | 0.67 | 0.67 | 0.67 | 0.67 | 0.45 | 3.15 | 0.81 | 1.6 | 0.67 | 0.67 | 0 | 0.22 | 0.67 | 0.09 | 0.9 | 0.45 | 0.67 | 0.67 | 0 | 1.93
| Analytical results for Mn in OREAS 610 (Certified Value 66 ppm) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 1
| 1 | 1 | 67.5 | 74 | 70 | 69 | 62 | 69 | 65 | 66 | 66 | 70 | 64.5 | 65.317 | 63 | 65 | 70 | 66.7128 | 63 | 66.362 | 64 | 70 | 63 | 65 | 90 | 78.01 | 1
| 1 | 2 | 59.1 | 72 | 68 | 69 | 63 | 68 | 65 | 67 | 63 | 71 | 63.5 | 65.7509 | 63 | 63 | 71 | 66.0613 | 63 | 64.062 | 65 | 73 | 64 | 64 | 90 | 79.22 | 32
| 1 | 3 | 59.8 | 73 | 69 | 65 | 62 | 67 | 65 | 67 | 65 | 72 | 59.2 | 64.0826 | 62 | 64 | 69 | 67.0099 | 64 | 65.555 | 64 | 70 | 64 | 65 | 91 | 78.6 | 16
| 1 | 4 | 67.6 | 72 | 71 | 67 | 62 | 69 | 65 | 66 | 64 | 70 | 58.2 | 64.6271 | 64 | 64 | 69 | 66.6704 | 61 | 66.146 | 64 | 72 | 64 | 65 | 89 | 78.64 | 65.5196
| 1 | 5 | 59.6 | 74 | 72 | 65 | 62 | 68 | 64 | 63 | 68 | 71 | 58.3 | 63.9301 | 63 | 65 | 67 | 66.7909 | 62 | 64.843 | 66 | 67 | 63 | 64 | 89 | 76.85 | 90
| 1 | 6 | 65.7 | 74 | 69 | 65 | 62 | 68 | 65 | 64 | 65 | 71 | 62.1 | 66.4149 | 64 | 64 | 66 | 66.8699 | 62 | 65.162 | 64 | 72 | 62 | 64 | 89 | 77.06
| Mean |  | 63.2167 | 73.1667 | 69.8333 | 66.6667 | 62.1667 | 68.1667 | 64.8333 | 65.5 | 65.1667 | 70.8333 | 60.9667 | 65.0204 | 63.1667 | 64.1667 | 68.6667 | 66.6859 | 62.5 | 65.355 | 64.5 | 70.6667 | 63.3333 | 64.5 | 89.6667 | 78.0633
| Median |  | 62.75 | 73.5 | 69.5 | 66 | 62 | 68 | 65 | 66 | 65 | 71 | 60.65 | 64.9721 | 63 | 64 | 69 | 66.7518 | 62.5 | 65.3585 | 64 | 71 | 63.5 | 64.5 | 89.5 | 78.305
| Std Dev. |  | 4.1335 | 0.98319 | 1.472 | 1.9664 | 0.40825 | 0.75277 | 0.40825 | 1.6432 | 1.7224 | 0.75277 | 2.7595 | 0.97869 | 0.75277 | 0.75277 | 1.8619 | 0.32891 | 1.0488 | 0.85451 | 0.83666 | 2.1602 | 0.8165 | 0.54772 | 0.8165 | 0.94235
| Rel.Std.Dev. |  | 6.539% | 1.344% | 2.108% | 2.95% | 0.6567% | 1.104% | 0.6297% | 2.509% | 2.643% | 1.063% | 4.526% | 1.505% | 1.192% | 1.173% | 2.712% | 0.4932% | 1.678% | 1.307% | 1.297% | 3.057% | 1.289% | 0.8492% | 0.9106% | 1.207%
| PDM3 |  | -3.515% | 11.67% | 6.584% | 1.751% | -5.117% | 4.04% | -1.047% | -0.02989% | -0.5386% | 8.11% | -6.949% | -0.7618% | -3.591% | -2.065% | 4.803% | 1.78% | -4.609% | -0.2512% | -1.556% | 7.856% | -3.337% | -1.556% | 36.85% | 19.15%
| Z-Score (Absolute) |  | 0.67 | 2.58 | 1.49 | 0.46 | 1.01 | 0.95 | 0.14 | 0.08 | 0.03 | 1.82 | 1.4 | 0.08 | 0.68 | 0.36 | 1.11 | 0.46 | 0.9 | 0.03 | 0.25 | 1.76 | 0.63 | 0.25 | 7.95 | 4.17
| Analytical results for Mo in OREAS 610 (Certified Value 4.47 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 4.4 | 4.83 | 4.74 | 4.48 | 4.53 | 4.45 | 4.57 | 4.48 | 4.3 | 4.99 | 4.7096 | 5 | 4.17 | 5.18 | 5.5409 | 4.61 | 4.286 | 4.88 | 4.36 | 4 | 4.4 | 5 | 6.834 |  | 1
| 1 | 2 | 4.3 | 4.65 | 4.56 | 4.19 | 4.41 | 4.75 | 4.5 | 4.55 | 4.3 | 4.89 | 4.6712 | 5 | 3.66 | 5.09 | 4.8208 | 4.17 | 3.985 | 4.88 | 4.27 | 4 | 4.5 | 5 | 4.843 |  | 33
| 1 | 3 | 4.2 | 4.58 | 4.36 | 4.6 | 4.61 | 4.61 | 4.59 | 4.33 | 4.4 | 4.8 | 4.4975 | 4 | 4.27 | 5.51 | 4.8439 | 4.22 | 3.775 | 4.9 | 4.48 | 4 | 4.4 | 4 | 4.852 |  | 16
| 1 | 4 | 4.4 | 4.73 | 4.65 | 4.31 | 4.52 | 4.69 | 4.39 | 4.43 | 4.3 | 4.93 | 4.6484 | 5 | 3.84 | 5.3 | 5.3887 | 4.31 | 4.252 | 4.77 | 4.18 | 4 | 4.7 | 5 | 5.139 |  | 4.4737
| 1 | 5 | 4.4 | 4.55 | 4.15 | 4.33 | 5.01 | 4.4 | 4.3 | 4.65 | 4.4 | 4.69 | 4.5888 | 5 | 4 | 5.26 | 4.5343 | 3.68 | 3.835 | 4.83 | 4.1 | 4 | 4.9 | 4 | 6.13 |  | 91
| 1 | 6 | 4.4 | 4.75 | 4.63 | 4.11 | 4.58 | 4.79 | 4.51 | 4.75 | 4.4 | 4.71 | 4.5046 | 4 | 3.89 | 5.15 | 4.7117 | 4.1 | 3.861 | 4.99 | 4.2 | 4 | 4.4 | 5 | 7.121
| Mean |  | 4.35 | 4.6817 | 4.515 | 4.3367 | 4.61 | 4.615 | 4.4767 | 4.5317 | 4.35 | 4.835 | 4.6033 | 4.6667 | 3.9717 | 5.2483 | 4.9734 | 4.1817 | 3.999 | 4.875 | 4.265 | 4 | 4.55 | 4.6667 | 5.8198
| Median |  | 4.4 | 4.69 | 4.595 | 4.32 | 4.555 | 4.65 | 4.505 | 4.515 | 4.35 | 4.845 | 4.6186 | 5 | 3.945 | 5.22 | 4.8323 | 4.195 | 3.923 | 4.88 | 4.235 | 4 | 4.45 | 5 | 5.6345
| Std Dev. |  | 0.083666 | 0.10741 | 0.2197 | 0.18085 | 0.20756 | 0.15997 | 0.1113 | 0.15211 | 0.054772 | 0.12161 | 0.088426 | 0.5164 | 0.22373 | 0.14878 | 0.39898 | 0.30328 | 0.22031 | 0.073417 | 0.13708 | 0 | 0.20736 | 0.5164 | 1.0171
| Rel.Std.Dev. |  | 1.923% | 2.294% | 4.866% | 4.17% | 4.502% | 3.466% | 2.486% | 3.357% | 1.259% | 2.515% | 1.921% | 11.07% | 5.633% | 2.835% | 8.022% | 7.253% | 5.509% | 1.506% | 3.214% | 0% | 4.557% | 11.07% | 17.48%
| PDM3 |  | -2.766% | 4.648% | 0.9224% | -3.064% | 3.046% | 3.158% | 0.06556% | 1.295% | -2.766% | 8.075% | 2.897% | 4.313% | -11.22% | 17.31% | 11.17% | -6.528% | -10.61% | 8.969% | -4.666% | -10.59% | 1.705% | 4.313% | 30.09%
| Z-Score (Absolute) |  | 0.65 | 0.48 | 0.09 | 0.7 | 0.24 | 0.25 | 0.22 | 0.03 | 0.65 | 1 | 0.21 | NA | 1.94 | 2.42 | 1.48 | 1.23 | 1.85 | 1.14 | 0.94 | NA | 0.03 | NA | 4.37
Results from laboratories 13, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 610 (Certified Value 0.049 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.06 | 0.045 | 0.078 | 0.05 | 0.05 | 0.04 | 0.05 | 0.06 | 0.06 | 0.047 | 0.041725 | 0.05 | 0.048 | 0.04 | 0.062923 | 0.03 | 0.044 | 0.04 | 0.05 | 0.05 | 0.09 | 0.05 | 0.1391 |  | 1
| 1 | 2 | 0.06 | 0.046 | 0.074 | 0.05 | 0.05 | 0.04 | 0.05 | 0.06 | 0.06 | 0.046 | 0.044 | 0.04 | 0.045 | 0.04 | 0.064526 | 0.03 | 0.045 | 0.05 | 0.05 | 0.05 | 0.09 | 0.05 | 0.13964 |  | #N/A
| 1 | 3 | 0.06 | 0.046 | 0.068 | 0.05 | 0.05 | 0.04 | 0.05 | 0.06 | 0.06 | 0.04 | 0.040164 | 0.04 | 0.047 | 0.03 | 0.067681 | 0.03 | 0.044 | 0.06 | 0.05 | 0.05 | 0.09 | 0.05 | 0.1368 |  | 16
| 1 | 4 | 0.06 | 0.046 | 0.068 | 0.05 | 0.05 | 0.04 | 0.05 | 0.06 | 0.06 | 0.04 | 0.043747 | 0.04 | 0.047 | 0.03 | 0.060898 | 0.03 | 0.044 | 0.05 | 0.05 | 0.05 | 0.08 | 0.05 | 0.1403 |  | 0.048667
| 1 | 5 | 0.07 | 0.046 | 0.066 | 0.05 | 0.05 | 0.04 | 0.04 | 0.06 | 0.06 | 0.039 | 0.040331 | 0.04 | 0.048 | 0.03 | 0.058546 | 0.03 | 0.044 | 0.05 | 0.04 | 0.05 | 0.08 | 0.05 | 0.13174 |  | 92
| 1 | 6 | 0.07 | 0.046 | 0.068 | 0.05 | 0.05 | 0.04 | 0.04 | 0.06 | 0.06 | 0.043 | 0.04306 | 0.04 | 0.047 | 0.03 | 0.055654 | 0.03 | 0.043 | 0.04 | 0.04 | 0.05 | 0.08 | 0.05 | 0.12445
| Mean |  | 0.063333 | 0.045833 | 0.070333 | 0.05 | 0.05 | 0.04 | 0.046667 | 0.06 | 0.06 | 0.0425 | 0.042171 | 0.041667 | 0.047 | 0.033333 | 0.061705 | 0.03 | 0.044 | 0.048333 | 0.046667 | 0.05 | 0.085 | 0.05 | 0.13534
| Median |  | 0.06 | 0.046 | 0.068 | 0.05 | 0.05 | 0.04 | 0.05 | 0.06 | 0.06 | 0.0415 | 0.042393 | 0.04 | 0.047 | 0.03 | 0.06191 | 0.03 | 0.044 | 0.05 | 0.05 | 0.05 | 0.085 | 0.05 | 0.13795
| Std Dev. |  | 0.005164 | 0.00040825 | 0.0046332 | 7.6012e-18 | 7.6012e-18 | 0 | 0.005164 | 0 | 0 | 0.0033912 | 0.0016872 | 0.0040825 | 0.0010954 | 0.005164 | 0.0042987 | 0 | 0.00063246 | 0.0075277 | 0.005164 | 7.6012e-18 | 0.0054772 | 7.6012e-18 | 0.0061791
| Rel.Std.Dev. |  | 8.154% | 0.8907% | 6.588% | 1.52e-14% | 1.52e-14% | 0% | 11.07% | 0% | 0% | 7.979% | 4.001% | 9.798% | 2.331% | 15.49% | 6.967% | 0% | 1.437% | 15.57% | 11.07% | 1.52e-14% | 6.444% | 1.52e-14% | 4.566%
| PDM3 |  | 30.14% | -5.823% | 44.52% | 2.739% | 2.739% | -17.81% | -4.11% | 23.29% | 23.29% | -12.67% | -13.35% | -14.38% | -3.426% | -31.51% | 26.79% | -38.36% | -9.59% | -0.6858% | -4.11% | 2.739% | 74.66% | 2.739% | 178.1%
| Z-Score (Absolute) |  | 1.64 | 0.27 | 2.41 | 0.18 | 0.18 | 0.91 | 0.18 | 1.28 | 1.28 | 0.64 | 0.67 | 0.73 | 0.15 | 1.64 | 1.46 | 2.01 | 0.47 | 0 | 0.18 | 0.18 | 4.01 | 0.18 | 9.52
| Analytical results for Nb in OREAS 610 (Certified Value 0.16 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [7 empty] | 2
| 1 | 1 | < 0.2 | 0.07 | 0.3 | 0.16 | 0.18 | 0.13 | 0.17 | 0.18 | 0.4 | 0.12 | < 0.2 | 0.32 | < 1 | < 0.2 | 0.21 | 0.17 | 0.2 | [7 empty] | 1
| 1 | 2 | < 0.2 | 0.06 | 0.3 | 0.16 | 0.18 | 0.14 | 0.17 | 0.19 | 0.4 | 0.13 | < 0.2 | 0.34 | < 1 | < 0.2 | 0.22 | 0.15 | 0.2 | [7 empty] | 34
| 1 | 3 | < 0.2 | 0.08 | 0.2 | 0.17 | 0.16 | 0.13 | 0.16 | 0.18 | 0.4 | 0.12 | < 0.2 | 0.34 | < 1 | < 0.2 | 0.21 | 0.17 | 0.1 | [7 empty] | 16
| 1 | 4 | 0.2 | 0.06 | 0.3 | 0.16 | 0.16 | 0.13 | 0.17 | 0.18 | 0.4 | 0.12 | < 0.2 | 0.34 | < 1 | < 0.2 | 0.21 | 0.14 | 0.1 | [7 empty] | 0.16208
| 1 | 5 | < 0.2 | 0.07 | 0.2 | 0.15 | 0.18 | 0.13 | 0.17 | 0.18 | 0.4 | 0.12 | < 0.2 | 0.37 | < 1 | < 0.2 | 0.2 | 0.14 | 0.1 | [7 empty] | 93
| 1 | 6 | < 0.2 | 0.06 | 0.3 | 0.16 | 0.17 | 0.14 | 0.15 | 0.19 | 0.4 | 0.13 | < 0.2 | 0.34 | < 1 | < 0.2 | 0.21 | 0.13 | 0.1
| Mean |  | 0.2 | 0.066667 | 0.26667 | 0.16 | 0.17167 | 0.13333 | 0.165 | 0.18333 | 0.4 | 0.12333 |  | 0.34167 |  |  | 0.21 | 0.15 | 0.13333
| Median |  | 0.2 | 0.065 | 0.3 | 0.16 | 0.175 | 0.13 | 0.17 | 0.18 | 0.4 | 0.12 |  | 0.34 |  |  | 0.21 | 0.145 | 0.1
| Std Dev. |  |  | 0.008165 | 0.05164 | 0.0063246 | 0.0098319 | 0.005164 | 0.0083666 | 0.005164 | 6.0809e-17 | 0.005164 |  | 0.016021 |  |  | 0.0063246 | 0.016733 | 0.05164
| Rel.Std.Dev. |  |  | 12.25% | 19.36% | 3.953% | 5.727% | 3.873% | 5.071% | 2.817% | 1.52e-14% | 4.187% |  | 4.689% |  |  | 3.012% | 11.16% | 38.73%
| PDM3 |  | 23.39% | -58.87% | 64.52% | -1.285% | 5.913% | -17.74% | 1.799% | 13.11% | 146.8% | -23.91% |  | 110.8% |  |  | 29.56% | -7.455% | -17.74%
| Z-Score (Absolute) |  | NA | 1.72 | NA | 0 | 0.21 | 0.49 | 0.09 | 0.43 | NA | 0.67 | 1.1 | 3.34 | 6.25 | 1.1 | 0.92 | 0.18 | NA
Results from laboratories 2, 4, 10 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 610 (Indicative Value 7.89 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | 2.02: Lab | [7 empty] | 1
| No | No | 2.02: 2.04 | [7 empty] | ppm
|  |  | 2.02: AR*MS | [7 empty] | 2
|  |  | 2.02: 0.5g | [7 empty] | 2
| 1 | 1 | 2.02: 7.95 | [7 empty] | 1
| 1 | 2 | 2.02: 8.28 | [7 empty] | 10
| 1 | 3 | 2.02: 8.08 | [7 empty] | 16
| 1 | 4 | 2.02: 7.78 | [7 empty] | 7.885
| 1 | 5 | 2.02: 7.33 | [7 empty] | 16
| 1 | 6 | 2.02: 7.89
| Mean |  | 2.02: 7.885
| Median |  | 2.02: 7.92
| Std Dev. |  | 2.02: 0.32167
| Rel.Std.Dev. |  | 2.02: 4.079%
| PDM3 |  | 2.02: 2.22e-14%
Z-Score (Absolute) | bar: 2.02
| Analytical results for Ni in OREAS 610 (Certified Value 24.3 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 1
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 2
| 1 | 1 | 21.4 | 26 | 26 | 26.1 | 23.8 | 24.3 | 23.7 | 24.1 | 24.5 | 25 | 23.7 | 24.3291 | 25 | 23 | 25.7 | 22.348 | 23.6 | 25.814 | 22.2 | 26.6 | 23 | 28.2 | 26 | 30.98 | 1
| 1 | 2 | 18.1 | 25 | 26.5 | 25.4 | 23.8 | 24.4 | 24.6 | 24.1 | 23.4 | 25 | 23.4 | 24.5278 | 24 | 21.4 | 26.2 | 21.8311 | 23.4 | 24.952 | 23.1 | 26.7 | 23 | 30 | 26 | 29.19 | 21
| 1 | 3 | 20.4 | 25 | 25.5 | 22.8 | 22.6 | 24.8 | 23.9 | 23.5 | 24.5 | 26 | 22.6 | 23.5939 | 24 | 22.3 | 25.8 | 21.8738 | 23.7 | 25.498 | 23 | 26.3 | 23 | 28.8 | 26 | 23.4 | 16
| 1 | 4 | 21 | 25 | 26.6 | 24.3 | 23.5 | 24 | 24.9 | 22.9 | 24.4 | 25 | 22.9 | 24.6659 | 24 | 20.1 | 25.9 | 21.9813 | 24 | 25.399 | 22.6 | 26.1 | 22 | 30.3 | 25 | 27.61 | 24.2843
| 1 | 5 | 17.4 | 26 | 25.9 | 23.1 | 23.1 | 23.3 | 24.2 | 23.1 | 25.5 | 26 | 22.6 | 24.021 | 25 | 22 | 25.4 | 21.8761 | 23.7 | 25.391 | 23.5 | 24.2 | 23 | 30.4 | 25 | 24.82 | 94
| 1 | 6 | 18.8 | 25 | 26 | 24.2 | 22.8 | 24.6 | 24.2 | 21.9 | 25.1 | 26 | 23 | 23.9386 | 25 | 21.5 | 25.9 | 21.7522 | 23.2 | 25.629 | 22.9 | 26.9 | 23 | 30.2 | 25 | 26.03
| Mean |  | 19.5167 | 25.3333 | 26.0833 | 24.3167 | 23.2667 | 24.2333 | 24.25 | 23.2667 | 24.5667 | 25.5 | 23.0333 | 24.1794 | 24.5 | 21.7167 | 25.8167 | 21.9438 | 23.6 | 25.4472 | 22.8833 | 26.1333 | 22.8333 | 29.65 | 25.5 | 27.005
| Median |  | 19.6 | 25 | 26 | 24.25 | 23.3 | 24.35 | 24.2 | 23.3 | 24.5 | 25.5 | 22.95 | 24.1751 | 24.5 | 21.75 | 25.85 | 21.8749 | 23.65 | 25.4485 | 22.95 | 26.45 | 23 | 30.1 | 25.5 | 26.82
| Std Dev. |  | 1.6449 | 0.5164 | 0.40702 | 1.2766 | 0.51251 | 0.53166 | 0.44159 | 0.83347 | 0.71461 | 0.54772 | 0.44121 | 0.40157 | 0.54772 | 0.98268 | 0.26394 | 0.21151 | 0.27568 | 0.29013 | 0.4446 | 0.98928 | 0.40825 | 0.92033 | 0.54772 | 2.8164
| Rel.Std.Dev. |  | 8.428% | 2.038% | 1.56% | 5.25% | 2.203% | 2.194% | 1.821% | 3.582% | 2.909% | 2.148% | 1.916% | 1.661% | 2.236% | 4.525% | 1.022% | 0.9639% | 1.168% | 1.14% | 1.943% | 3.785% | 1.788% | 3.104% | 2.148% | 10.43%
| PDM3 |  | -19.63% | 4.32% | 7.408% | 0.1331% | -4.191% | -0.21% | -0.1414% | -4.191% | 1.163% | 5.006% | -5.151% | -0.4321% | 0.8881% | -10.57% | 6.31% | -9.638% | -2.818% | 4.788% | -5.769% | 7.614% | -5.975% | 22.1% | 5.006% | 11.2%
| Z-Score (Absolute) |  | 2.64 | 0.58 | 1 | 0.02 | 0.56 | 0.03 | 0.02 | 0.56 | 0.16 | 0.67 | 0.69 | 0.06 | 0.12 | 1.42 | 0.85 | 1.3 | 0.38 | 0.65 | 0.78 | 1.03 | 0.8 | 2.97 | 0.67 | 1.51
| Analytical results for P in OREAS 610 (Certified Value 249 ppm) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 264 | 240 | 300 | 240 | 260 | 240 | 240 | 250 | 257 | 248 | 521.8751 | 190 | 250 | 270 | 247.3441 | 400 | 420 | 245 | 300 | 300 | 310 | 350 | 265.2 |  | 1
| 1 | 2 | 250 | 250 | 300 | 240 | 250 | 240 | 240 | 240 | 266 | 236 | 535.8996 | 190 | 240 | 260 | 247.2157 | 400 | 400 | 253 | 300 | 300 | 300 | 350 | 239.2 |  | #N/A
| 1 | 3 | 263 | 240 | 300 | 240 | 250 | 240 | 240 | 240 | 268 | 202 | 544.6976 | 190 | 240 | 260 | 248.6777 | 400 | 410 | 255 | 300 | 300 | 310 | 350 | 233.2 |  | 16
| 1 | 4 | 259 | 250 | 290 | 240 | 250 | 240 | 240 | 240 | 260 | 200 | 558.3997 | 190 | 250 | 270 | 244.9149 | 400 | 420 | 251 | 300 | 300 | 300 | 340 | 217.2 |  | 249.4704
| 1 | 5 | 261 | 240 | 300 | 240 | 250 | 240 | 230 | 260 | 267 | 206 | 524.4869 | 190 | 250 | 260 | 243.1286 | 400 | 420 | 250 | 300 | 300 | 300 | 340 | 211.2 |  | 95
| 1 | 6 | 261 | 250 | 300 | 230 | 250 | 240 | 240 | 250 | 264 | 224 | 544.4198 | 190 | 250 | 260 | 233.2455 | 300 | 420 | 252 | 300 | 300 | 300 | 340 | 225.2
| Mean |  | 259.6667 | 245 | 298.3333 | 238.3333 | 251.6667 | 240 | 238.3333 | 246.6667 | 263.6667 | 219.3333 | 538.2965 | 190 | 246.6667 | 263.3333 | 244.0877 | 383.3333 | 415 | 251 | 300 | 300 | 303.3333 | 345 | 231.8667
| Median |  | 261 | 245 | 300 | 240 | 250 | 240 | 240 | 245 | 265 | 215 | 540.1597 | 190 | 250 | 260 | 246.0653 | 400 | 420 | 251.5 | 300 | 300 | 300 | 345 | 229.2
| Std Dev. |  | 5.0465 | 5.4772 | 4.0825 | 4.0825 | 4.0825 | 0 | 4.0825 | 8.165 | 4.3205 | 19.8662 | 13.7748 | 0 | 5.164 | 5.164 | 5.6677 | 40.8248 | 8.3666 | 3.4059 | 0 | 0 | 5.164 | 5.4772 | 19.2527
| Rel.Std.Dev. |  | 1.943% | 2.236% | 1.368% | 1.713% | 1.622% | 0% | 1.713% | 3.31% | 1.639% | 9.058% | 2.559% | 0% | 2.094% | 1.961% | 2.322% | 10.65% | 2.016% | 1.357% | 0% | 0% | 1.702% | 1.588% | 8.303%
| PDM3 |  | 4.087% | -1.792% | 19.59% | -4.464% | 0.8804% | -3.796% | -4.464% | -1.124% | 5.691% | -12.08% | 115.8% | -23.84% | -1.124% | 5.557% | -2.158% | 53.66% | 66.35% | 0.6131% | 20.25% | 20.25% | 21.59% | 38.29% | -7.056%
| Z-Score (Absolute) |  | 0.58 | 0.2 | 2.64 | 0.56 | 0.15 | 0.47 | 0.56 | 0.12 | 0.79 | 1.57 | 15.41 | 3.13 | 0.12 | 0.77 | 0.25 | NA | 8.85 | 0.12 | NA | NA | 2.9 | 5.12 | 0.9
Results from laboratories 18, 21 & 22 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 610 (Certified Value 512 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 0
| 1 | 1 | 505 | 542 | 519.47 | 499 | 497 | 523 | 495 | 521 | 508 | 600 | 522.7 | 522.281 | 506 | 527.12 | 510 | 494.9445 | 517 | 535.385 | 526.4 | 486 | 473 | 546.7 | 550 | 607.9 | 1
| 1 | 2 | 495 | 513 | 521.92 | 484 | 504 | 525 | 491 | 521 | 482 | 607 | 516.4 | 531.0912 | 494 | 512.26 | 510 | 490.2715 | 522 | 533.017 | 517.2 | 495 | 477 | 536.7 | 550 | 594 | 23
| 1 | 3 | 481 | 535 | 513.96 | 474 | 497 | 517 | 486 | 532 | 498 | 613 | 532.2 | 525.6724 | 491 | 525.01 | 523 | 499.9542 | 527.2 | 529.804 | 516.2 | 481 | 479 | 534.9 | 550 | 580.1 | 16
| 1 | 4 | 504 | 537 | 515.46 | 479 | 489 | 522 | 491 | 521 | 499 | 603 | 527.3 | 534.6778 | 503 | 528.34 | 510 | 495.1469 | 513 | 532.958 | 521.6 | 473 | 462 | 543.4 | 510 | 586.2 | 511.6994
| 1 | 5 | 497 | 541 | 519.36 | 461 | 500 | 518 | 489 | 501 | 526 | 605 | 526.4 | 521.0802 | 503 | 533.82 | 510 | 498.489 | 518.6 | 530.374 | 520.9 | 442 | 474 | 540 | 530 | 582.3 | 96
| 1 | 6 | 509 | 535 | 524.93 | 482 | 488 | 518 | 495 | 487 | 509 | 612 | 513.5 | 532.0207 | 507 | 524.61 | 520 | 489.613 | 517.4 | 530.379 | 526 | 483 | 470 | 543.4 | 530 | 578.4
| Mean |  | 498.5 | 533.8333 | 519.1833 | 479.8333 | 495.8333 | 520.5 | 491.1667 | 513.8333 | 503.6667 | 606.6667 | 523.0833 | 527.8039 | 500.6667 | 525.1933 | 513.8333 | 494.7365 | 519.2 | 531.9862 | 521.3833 | 476.6667 | 472.5 | 540.85 | 536.6667 | 588.15
| Median |  | 500.5 | 536 | 519.415 | 480.5 | 497 | 520 | 491 | 521 | 503.5 | 606 | 524.55 | 528.3818 | 503 | 526.065 | 510 | 495.0457 | 518 | 531.6685 | 521.25 | 482 | 473.5 | 541.7 | 540 | 584.25
| Std Dev. |  | 10.0349 | 10.6286 | 4.0441 | 12.4807 | 6.2423 | 3.2711 | 3.4881 | 16.5459 | 14.6242 | 5.0859 | 7.0505 | 5.5873 | 6.5929 | 7.1509 | 6.0139 | 4.1882 | 4.8728 | 2.1681 | 4.2692 | 18.4246 | 6.0249 | 4.485 | 16.3299 | 11.155
| Rel.Std.Dev. |  | 2.013% | 1.991% | 0.7789% | 2.601% | 1.259% | 0.6285% | 0.7102% | 3.22% | 2.904% | 0.8383% | 1.348% | 1.059% | 1.317% | 1.362% | 1.17% | 0.8465% | 0.9385% | 0.4075% | 0.8188% | 3.865% | 1.275% | 0.8292% | 3.043% | 1.897%
| PDM3 |  | -2.58% | 4.326% | 1.463% | -6.227% | -3.101% | 1.72% | -4.013% | 0.417% | -1.57% | 18.56% | 2.225% | 3.147% | -2.156% | 2.637% | 0.417% | -3.315% | 1.466% | 3.965% | 1.893% | -6.846% | -7.661% | 5.697% | 4.879% | 14.94%
| Z-Score (Absolute) |  | 0.85 | 0.6 | 0 | 1.61 | 0.95 | 0.05 | 1.15 | 0.22 | 0.63 | 3.57 | 0.16 | 0.35 | 0.76 | 0.25 | 0.22 | 1 | 0 | 0.52 | 0.09 | 1.74 | 1.91 | 0.89 | 0.71 | 2.82
| Analytical results for Pd in OREAS 610 (Indicative Value 174 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.03; 2.02: 2.11; 2.03: 2.12 | ppb
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 0
|  |  | 2.01: 15g; 2.02: 0.5g; 2.03: 01g | 0
| 1 | 1 | 2.01: < 10; 2.02: 177; 2.03: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.02: 172; 2.03: < 10 | 1
| 1 | 3 | 2.01: < 10; 2.02: 172; 2.03: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.02: 171; 2.03: < 10 | 173.6667
| 1 | 5 | 2.01: < 10; 2.02: 178; 2.03: < 10 | 17
| 1 | 6 | 2.01: < 10; 2.02: 172; 2.03: < 10
| Mean |  | 2.02: 173.6667
| Median |  | 2.02: 172
| Std Dev. |  | 2.02: 3.0111
| Rel.Std.Dev. |  | 2.02: 1.734%
| PDM3 |  | 2.02: -1.998e-13%
Z-Score (Absolute) | bar: 2.01 to 2.03
| Analytical results for Pr in OREAS 610 (Indicative Value 1.93 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 2 | 1
| 1 | 2 | 2.01: 2 | 12
| 1 | 3 | 2.01: 1.9 | 16
| 1 | 4 | 2.01: 2 | 1.9333
| 1 | 5 | 2.01: 1.8 | 18
| 1 | 6 | 2.01: 1.9
| Mean |  | 2.01: 1.9333
| Median |  | 2.01: 1.95
| Std Dev. |  | 2.01: 0.08165
| Rel.Std.Dev. |  | 2.01: 4.223%
| PDM3 |  | 2.01: 1.776e-13%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Pt in OREAS 610 (Indicative Value < 2 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.03; 2.02: 2.11; 2.03: 2.12 | ppb
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 15g; 2.02: 0.5g; 2.03: 01g | 2
| 1 | 1 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 1
| 1 | 2 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 1
| 1 | 3 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 16
| 1 | 4 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | < 2
| 1 | 5 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 19
| 1 | 6 | 2.01: < 2; 2.02: < 1; 2.03: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.01 to 2.03
| Analytical results for Rb in OREAS 610 (Certified Value 7.63 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 11.65 | 8.1 | 11.2 | 7.6 | 7.9 | 7.6 | 7.8 | 8.3 | 11.5 | 8.26 | 11.7908 | 6.3 | 7.2 | 7.4 | 7.8 | 7.9 | [8 empty] | 1
| 1 | 2 | 11.79 | 8.4 | 11 | 7.5 | 7.8 | 7.9 | 7.9 | 8.2 | 11.9 | 8.25 | 12.1516 | 6 | 7.3 | 7.4 | 7.9 | 7.8 | [8 empty] | 39
| 1 | 3 | 11.1 | 8.2 | 10.4 | 7.4 | 7.8 | 7.9 | 7.1 | 8.4 | 11.7 | 7.85 | 11.8391 | 5.9 | 7.2 | 7.5 | 8 | 7.7 | [8 empty] | 16
| 1 | 4 | 11.98 | 8.2 | 10.5 | 7.5 | 7.6 | 8.1 | 7.8 | 8.5 | 11.6 | 7.91 | 12.2103 | 6.3 | 7.3 | 7.5 | 7.8 | 7.2 | [8 empty] | 7.6278
| 1 | 5 | 11.76 | 8.2 | 9.8 | 7.4 | 7.7 | 7.9 | 7.6 | 8.7 | 11.4 | 7.86 | 11.7632 | 6 | 7.3 | 7.6 | 7.1 | 7.1 | [8 empty] | 97
| 1 | 6 | 11.77 | 8.3 | 10.6 | 7.5 | 7.8 | 8 | 7.2 | 8.6 | 11.6 | 7.99 | 12.115 | 6 | 6.8 | 7.6 | 7.6 | 7
| Mean |  | 11.675 | 8.2333 | 10.5833 | 7.4833 | 7.7667 | 7.9 | 7.5667 | 8.45 | 11.6167 | 8.02 | 11.9783 | 6.0833 | 7.1833 | 7.5 | 7.7 | 7.45
| Median |  | 11.765 | 8.2 | 10.55 | 7.5 | 7.8 | 7.9 | 7.7 | 8.45 | 11.6 | 7.95 | 11.9771 | 6 | 7.25 | 7.5 | 7.8 | 7.45
| Std Dev. |  | 0.30125 | 0.10328 | 0.4916 | 0.075277 | 0.10328 | 0.16733 | 0.33862 | 0.18708 | 0.17224 | 0.18868 | 0.20167 | 0.17224 | 0.19408 | 0.089443 | 0.32249 | 0.3937
| Rel.Std.Dev. |  | 2.58% | 1.254% | 4.645% | 1.006% | 1.33% | 2.118% | 4.475% | 2.214% | 1.483% | 2.353% | 1.684% | 2.831% | 2.702% | 1.193% | 4.188% | 5.285%
| PDM3 |  | 53.06% | 7.939% | 38.75% | -1.894% | 1.821% | 3.569% | -0.8012% | 10.78% | 52.29% | 5.142% | 57.04% | -20.25% | -5.827% | -1.675% | 0.9468% | -2.331%
| Z-Score (Absolute) |  | 6.61 | 0.69 | 4.73 | 0.6 | 0.11 | 0.11 | 0.46 | 1.06 | 6.51 | 0.32 | 7.14 | 3.01 | 1.12 | 0.57 | 0.23 | 0.66
| Analytical results for Re in OREAS 610 (Indicative Value 1.1 ppb) | [25 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.18 | 2.20 | 2.23 | [10 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.25g | 0.25g | [10 empty] | 2
| 1 | 1 | < 50 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | < 50 | 1 | < 50 | < 2 | 2 | < 5 | [10 empty] | 1
| 1 | 2 | < 50 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | < 50 | 1 | < 50 | < 2 | 3 | < 5 | [10 empty] | 14
| 1 | 3 | < 50 | < 1 | 1 | 1 | 1 | 1 | 1 | < 1 | < 50 | < 1 | < 50 | < 2 | 2 | < 5 | [10 empty] | 16
| 1 | 4 | < 50 | 2 | 1 | 2 | 2 | 1 | 1 | 1 | < 50 | < 1 | < 50 | < 2 | 2 | < 5 | [10 empty] | 1.1042
| 1 | 5 | < 50 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | < 50 | 1 | < 50 | < 2 | 2 | < 5 | [10 empty] | 20
| 1 | 6 | < 50 | 1 | 2 | 2 | 1 | 1 | < 1 | 1 | < 50 | < 1 | < 50 | < 2 | 2 | < 5
| Mean |  |  | 1.2 | 1.1667 | 1.6667 | 1.3333 | 1 | 1 | 1 |  | 1 |  |  | 2.1667
| Median |  |  | 1 | 1 | 2 | 1 | 1 | 1 | 1 |  | 1 |  |  | 2
| Std Dev. |  |  | 0.44721 | 0.40825 | 0.5164 | 0.5164 | 0 | 0 | 0 |  | 0 |  |  | 0.40825
| Rel.Std.Dev. |  |  | 37.27% | 34.99% | 30.98% | 38.73% | 0% | 0% | 0% |  | 0% |  |  | 18.84%
| PDM3 |  |  | 8.679% | 5.66% | 50.94% | 20.75% | -9.434% | -9.434% | -9.434% |  | -9.434% |  |  | 96.23%
| Z-Score (Absolute) |  | 42.71 | 0.3 | 0.15 | 0.75 | 0.15 | 0.45 | 0.6 | 0.6 | 42.71 | 0.9 | 42.71 | 0.45 | 1.65 | 2.25
| Analytical results for S in OREAS 610 (Certified Value 2.65 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 2.91 | 2.63 | 2.62 | 2.66 | 2.77 | 2.62 | 2.58 | 2.71 | 2.83 | > 5 | 2.998 | 2.65 | 2.6 | 2.77 | 2.577 | 2.64 | 2.666 | 2.57 | 2.54 | 2.74 | 3.22 | 3.73 | 2.694 |  | 1
| 1 | 2 | 2.78 | 2.67 | 2.657 | 2.66 | 2.77 | 2.61 | 2.59 | 2.6 | 2.83 | 2.48 | 3.0131 | 2.61 | 2.53 | 2.75 | 2.5953 | 2.64 | 2.7 | 2.64 | 2.62 | 2.76 | 3.19 | 3.76 | 2.707 |  | 16
| 1 | 3 | 2.84 | 2.66 | 2.676 | 2.66 | 2.74 | 2.57 | 2.63 | 2.68 | 2.84 | 2.37 | 2.9862 | 2.59 | 2.57 | 2.66 | 2.6465 | 2.64 | 2.69 | 2.57 | 2.62 | 2.79 | 3.23 | 3.74 | 2.62 |  | 16
| 1 | 4 | 2.8 | 2.66 | 2.573 | 2.62 | 2.75 | 2.59 | 2.56 | 2.69 | 2.81 | 2.41 | 2.9555 | 2.66 | 2.59 | 2.71 | 2.4564 | 2.65 | 2.778 | 2.59 | 2.58 | 2.73 | 3.2 | 3.65 | 2.733 |  | 2.6502
| 1 | 5 | 2.87 | 2.64 | 2.565 | 2.64 | 2.74 | 2.57 | 2.51 | 2.8 | 2.82 | 2.37 | 2.9546 | 2.65 | 2.59 | 2.61 | 2.4252 | 2.66 | 2.782 | 2.58 | 2.31 | 2.75 | 3.23 | 3.69 | 2.746 |  | 98
| 1 | 6 | 2.86 | 2.66 | 2.603 | 2.6 | 2.72 | 2.61 | 2.56 | 2.72 | 2.78 | 2.34 | 2.9894 | 2.66 | 2.55 | 2.58 | 2.4124 | 2.62 | 2.784 | 2.56 | 2.58 | 2.76 | 3.23 | 3.7 | 2.789
| Mean |  | 2.8433 | 2.6533 | 2.6157 | 2.64 | 2.7483 | 2.595 | 2.5717 | 2.7 | 2.8183 | 2.394 | 2.9828 | 2.6367 | 2.5717 | 2.68 | 2.5188 | 2.6417 | 2.7333 | 2.585 | 2.5417 | 2.755 | 3.2167 | 3.7117 | 2.7148
| Median |  | 2.85 | 2.66 | 2.6115 | 2.65 | 2.745 | 2.6 | 2.57 | 2.7 | 2.825 | 2.37 | 2.9878 | 2.65 | 2.58 | 2.685 | 2.5167 | 2.64 | 2.739 | 2.575 | 2.58 | 2.755 | 3.225 | 3.715 | 2.72
| Std Dev. |  | 0.04761 | 0.015055 | 0.044532 | 0.025298 | 0.019408 | 0.021679 | 0.039707 | 0.064807 | 0.02137 | 0.054129 | 0.023427 | 0.029439 | 0.027142 | 0.07642 | 0.099535 | 0.013292 | 0.053765 | 0.02881 | 0.11737 | 0.020736 | 0.017512 | 0.039707 | 0.05708
| Rel.Std.Dev. |  | 1.674% | 0.5674% | 1.702% | 0.9583% | 0.7062% | 0.8354% | 1.544% | 2.4% | 0.7582% | 2.261% | 0.7854% | 1.117% | 1.055% | 2.851% | 3.952% | 0.5032% | 1.967% | 1.114% | 4.618% | 0.7527% | 0.5444% | 1.07% | 2.103%
| PDM3 |  | 7.286% | 0.117% | -1.304% | -0.3861% | 3.702% | -2.084% | -2.964% | 1.878% | 6.343% | -9.668% | 12.55% | -0.5118% | -2.964% | 1.123% | -4.959% | -0.3232% | 3.136% | -2.461% | -4.096% | 3.953% | 21.37% | 40.05% | 2.438%
| Z-Score (Absolute) |  | 1.57 | 0 | 0.31 | 0.11 | 0.78 | 0.48 | 0.67 | 0.39 | 1.36 | 2.14 | 2.72 | 0.14 | 0.67 | 0.22 | 1.11 | 0.1 | 0.66 | 0.56 | 0.92 | 0.84 | 4.65 | 8.74 | 0.51
| Analytical results for Sb in OREAS 610 (Certified Value 265 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 210.71 | 264.49 | 282 | 251 | 282 | 259 | 272 | 278 | 241 | 276.54 | 269.1065 | 261 | 283.34 | 257 | 211.3378 | 183.3 | 243.579 | 223.43 | 261 | 211 | 289.8 | 270 | 265.9 |  | 1
| 1 | 2 | 209.46 | 265.19 | 270 | 255 | 286 | 253 | 275 | 270 | 239 | 273.67 | 273.4958 | 261 | 282.82 | 252 | 215.966 | 183.23 | 244.012 | 219.01 | 263 | 212 | 290.2 | 270 | 264.3 |  | 41
| 1 | 3 | 207.21 | 267.14 | 260 | 250 | 285 | 253 | 278 | 277 | 242 | 272.59 | 259.5307 | 251 | 281.38 | 245 | 221.9427 | 183.17 | 240.427 | 216.02 | 269 | 215 | 290 | 270 | 262.7 |  | 16
| 1 | 4 | 215.37 | 264.66 | 266 | 249 | 283 | 254 | 272 | 277 | 237 | 275.61 | 268.3354 | 261 | 281.01 | 255 | 202.3384 | 179.9 | 243.853 | 216.06 | 260 | 209 | 292.6 | 260 | 261.1 |  | 264.7681
| 1 | 5 | 212.91 | 267.49 | 254 | 252 | 286 | 251 | 268 | 287 | 244 | 273.55 | 258.3473 | 256 | 283.03 | 253 | 207.6103 | 183.39 | 238.812 | 218.77 | 238 | 211 | 287.8 | 260 | 259.5 |  | 99
| 1 | 6 | 218.06 | 267.29 | 263 | 249 | 282 | 258 | 270 | 282 | 248 | 267.12 | 263.0531 | 267 | 280.57 | 243 | 198.7585 | 184.31 | 240.912 | 219.08 | 258 | 223 | 287 | 260 | 257.9
| Mean |  | 212.2867 | 266.0433 | 265.8333 | 251 | 284 | 254.6667 | 272.5 | 278.5 | 241.8333 | 273.18 | 265.3115 | 259.5 | 282.025 | 250.8333 | 209.659 | 182.8833 | 241.9325 | 218.7283 | 258.1667 | 213.5 | 289.5667 | 265 | 261.9
| Median |  | 211.81 | 266.165 | 264.5 | 250.5 | 284 | 253.5 | 272 | 277.5 | 241.5 | 273.61 | 265.6942 | 261 | 282.1 | 252.5 | 209.4741 | 183.265 | 242.2455 | 218.89 | 260.5 | 211.5 | 289.9 | 265 | 261.9
| Std Dev. |  | 3.9869 | 1.4074 | 9.6003 | 2.2804 | 1.8974 | 3.1411 | 3.5637 | 5.6833 | 3.8687 | 3.3058 | 5.9589 | 5.4314 | 1.1777 | 5.6006 | 8.6034 | 1.5211 | 2.1803 | 2.7117 | 10.572 | 5.0498 | 1.9775 | 5.4772 | 2.9933
| Rel.Std.Dev. |  | 1.878% | 0.529% | 3.611% | 0.9085% | 0.6681% | 1.233% | 1.308% | 2.041% | 1.6% | 1.21% | 2.246% | 2.093% | 0.4176% | 2.233% | 4.104% | 0.8317% | 0.9012% | 1.24% | 4.095% | 2.365% | 0.6829% | 2.067% | 1.143%
| PDM3 |  | -19.82% | 0.4816% | 0.4023% | -5.2% | 7.264% | -3.815% | 2.92% | 5.186% | -8.662% | 3.177% | 0.2052% | -1.99% | 6.518% | -5.263% | -20.81% | -30.93% | -8.625% | -17.39% | -2.493% | -19.36% | 9.366% | 0.08759% | -1.083%
| Z-Score (Absolute) |  | 2.33 | 0.32 | 0.31 | 0.42 | 1.21 | 0.24 | 0.64 | 0.94 | 0.87 | 0.67 | 0.29 | 0 | 1.11 | 0.43 | 2.46 | 3.78 | 0.87 | 2.01 | 0.07 | 2.27 | 1.48 | 0.27 | 0.12
| Analytical results for Sc in OREAS 610 (Certified Value 0.84 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 01g |  |  | 2
| 1 | 1 | 1 | 1.4 | 0.9 | 0.9 | 0.9 | 0.8 | 0.9 | 0.9 | 1 | 0.87 | < 1 | < 5 | 0.9 | 0.8 | < 5 | 0.7 | 0.726 | 1.1 | 0.8 | < 0.5 | 1 | < 1 |  |  | 1
| 1 | 2 | < 1 | 1.6 | 1.2 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8 | 1 | 0.84 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.719 | 1.1 | 0.9 | < 0.5 | 1 | < 1 |  |  | 42
| 1 | 3 | < 1 | 1.7 | 0.7 | 1 | 0.9 | 0.9 | 0.8 | 0.9 | 1 | 0.78 | < 1 | < 5 | 0.9 | 0.8 | < 5 | 0.7 | 0.735 | 1.1 | 0.8 | < 0.5 | 1 | < 1 |  |  | 16
| 1 | 4 | < 1 | 1.5 | 1.1 | 1 | 0.8 | 0.9 | 0.8 | 0.9 | 1 | 0.78 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.719 | 1.1 | 0.8 | < 0.5 | 0.9 | < 1 |  |  | 0.84412
| 1 | 5 | < 1 | 1.4 | 1.1 | 1 | 0.8 | 0.9 | 0.8 | 0.9 | 1 | 0.75 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.726 | 1 | 0.7 | < 0.5 | 0.9 | < 1 |  |  | 100
| 1 | 6 | < 1 | 1.7 | 1 | 0.9 | 0.8 | 0.9 | 0.8 | 0.9 | 1 | 0.81 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.726 | 1 | 0.8 | < 0.5 | 0.9 | < 1
| Mean |  | 1 | 1.55 | 1 | 0.95 | 0.85 | 0.88333 | 0.83333 | 0.88333 | 1 | 0.805 |  |  | 0.83333 | 0.8 |  | 0.7 | 0.72517 | 1.0667 | 0.8 |  | 0.95
| Median |  | 1 | 1.55 | 1.05 | 0.95 | 0.85 | 0.9 | 0.8 | 0.9 | 1 | 0.795 |  |  | 0.8 | 0.8 |  | 0.7 | 0.726 | 1.1 | 0.8 |  | 0.95
| Std Dev. |  |  | 0.13784 | 0.17889 | 0.054772 | 0.054772 | 0.040825 | 0.05164 | 0.040825 | 0 | 0.044159 |  |  | 0.05164 | 1.2162e-16 |  | 1.2162e-16 | 0.0059133 | 0.05164 | 0.063246 |  | 0.054772
| Rel.Std.Dev. |  |  | 8.893% | 17.89% | 5.766% | 6.444% | 4.622% | 6.197% | 4.622% | 0% | 5.486% |  |  | 6.197% | 1.52e-14% |  | 1.737e-14% | 0.8154% | 4.841% | 7.906% |  | 5.766%
| PDM3 |  | 18.47% | 83.62% | 18.47% | 12.54% | 0.6971% | 4.646% | -1.277% | 4.646% | 18.47% | -4.634% |  |  | -1.277% | -5.226% |  | -17.07% | -14.09% | 26.37% | -5.226% |  | 12.54%
| Z-Score (Absolute) |  | NA | 4.25 | 0.95 | 0.65 | 0.05 | 0.25 | 0.05 | 0.25 | NA | 0.22 | 2.05 | 9.95 | 0.05 | 0.25 | 9.95 | 0.85 | 0.7 | 1.35 | 0.25 | 3.55 | 0.65 | 2.05
Results from laboratories 2 & 10 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 610 (Certified Value 27.7 ppm) | [25 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [5 empty] | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | [5 empty] | 1
| 1 | 1 | 26 | 29.1 | 36.4 | 30.4 | 32.7 | 31 | 30.7 | 28.6 | 26 | 28.39 | 27.1583 | 23.9 | 25.4051 | 20.1 | 29 | 26 | 30 | 19 | 27 | [5 empty] | 1
| 1 | 2 | 25 | 28.8 | 34.5 | 29 | 32.3 | 29.9 | 30.2 | 27.6 | 26 | 28.2 | 27.1729 | 21.6 | 23.6473 | 20.2 | 30 | 25 | 30 | 18.9 | 27 | [5 empty] | 43
| 1 | 3 | 24 | 28.5 | 33.6 | 29.4 | 32.8 | 29.8 | 29.2 | 28.9 | 26 | 28.76 | 26.8534 | 24.3 | 24.3256 | 19.3 | 30 | 25 | 30 | 19.4 | 27 | [5 empty] | 16
| 1 | 4 | 26 | 28.4 | 34.6 | 29.2 | 32.4 | 32.9 | 29 | 28.4 | 26 | 28.48 | 26.6138 | 20.7 | 24.8593 | 19.7 | 30 | 25 | 30 | 18.8 | 27 | [5 empty] | 27.7179
| 1 | 5 | 25 | 28.7 | 32.2 | 30 | 30.8 | 30.8 | 28.7 | 30 | 26 | 28.49 | 27.2644 | 22.2 | 23.6398 | 20.4 | 30 | 24 | 30 | 18.3 | 28 | [5 empty] | 101
| 1 | 6 | 26 | 29 | 34.8 | 29.6 | 32.7 | 32 | 31.4 | 31.1 | 26 | 28.19 | 27.5686 | 21.9 | 24.923 | 21.1 | 30 | 26 | 40 | 19.2 | 26
| Mean |  | 25.3333 | 28.75 | 34.35 | 29.6 | 32.2833 | 31.0667 | 29.8667 | 29.1 | 26 | 28.4183 | 27.1052 | 22.4333 | 24.4667 | 20.1333 | 29.8333 | 25.1667 | 31.6667 | 18.9333 | 27
| Median |  | 25.5 | 28.75 | 34.55 | 29.5 | 32.55 | 30.9 | 29.7 | 28.75 | 26 | 28.435 | 27.1656 | 22.05 | 24.5924 | 20.15 | 30 | 25 | 30 | 18.95 | 27
| Std Dev. |  | 0.8165 | 0.27386 | 1.391 | 0.52154 | 0.75211 | 1.2061 | 1.069 | 1.2522 | 0 | 0.21274 | 0.33236 | 1.3909 | 0.72361 | 0.61536 | 0.40825 | 0.75277 | 4.0825 | 0.37771 | 0.63246
| Rel.Std.Dev. |  | 3.223% | 0.9526% | 4.05% | 1.762% | 2.33% | 3.882% | 3.579% | 4.303% | 0% | 0.7486% | 1.226% | 6.2% | 2.958% | 3.056% | 1.368% | 2.991% | 12.89% | 1.995% | 2.342%
| PDM3 |  | -8.603% | 3.724% | 23.93% | 6.79% | 16.47% | 12.08% | 7.752% | 4.986% | -6.198% | 2.527% | -2.21% | -19.07% | -11.73% | -27.36% | 7.632% | -9.204% | 14.25% | -31.69% | -2.59%
| Z-Score (Absolute) |  | 0.72 | 0.29 | 1.96 | 0.55 | 1.35 | 0.98 | 0.63 | 0.4 | 0.52 | 0.2 | 0.2 | 1.59 | 0.98 | 2.27 | 0.62 | 0.77 | NA | 2.63 | 0.23
Results from laboratory 22 were removed due to their 10 ppm reading resolution.
| Analytical results for Sm in OREAS 610 (Indicative Value 1.55 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | 2.02: Lab | [5 empty] | 1
| No | No | 2.02: 2.04 | [5 empty] | ppm
|  |  | 2.02: AR*MS | [5 empty] | 2
|  |  | 2.02: 0.5g | [5 empty] | 2
| 1 | 1 | 2.02: 1.6 | [5 empty] | 1
| 1 | 2 | 2.02: 1.7 | [5 empty] | 8
| 1 | 3 | 2.02: 1.5 | [5 empty] | 16
| 1 | 4 | 2.02: 1.7 | [5 empty] | 1.55
| 1 | 5 | 2.02: 1.4 | [5 empty] | 14
| 1 | 6 | 2.02: 1.4
| Mean |  | 2.02: 1.55
| Median |  | 2.02: 1.55
| Std Dev. |  | 2.02: 0.13784
| Rel.Std.Dev. |  | 2.02: 8.893%
PDM3 | bar: 2.02
Z-Score (Absolute) | bar: 2.02
| Analytical results for Sn in OREAS 610 (Certified Value 24.8 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  |  | 2
| 1 | 1 | 24.4 | 25.7 | 25.3 | 24.8 | 25.9 | 25 | 26 | 24.6 | 21.5 | 26.08 | 23.9104 | 25.2 | 26.0071 | 21.7 | 24.805 | 26 | 24.4 | 24 | 26.3 | 24 | 21.44 |  |  |  | 1
| 1 | 2 | 23.8 | 27.3 | 24.4 | 24.1 | 26 | 25.5 | 25.6 | 23.5 | 21.8 | 25.86 | 24.368 | 25.6 | 25.3032 | 22 | 24.884 | 25.5 | 24 | 23 | 26 | 24 | 21.38 |  |  |  | 18
| 1 | 3 | 23.6 | 25.5 | 23.6 | 24.3 | 26 | 25.6 | 25.1 | 24.4 | 22 | 25.52 | 23.8434 | 24.8 | 25.5759 | 21.8 | 24.402 | 25.1 | 23.9 | 23 | 26.3 | 24 | 20.32 |  |  |  | 16
| 1 | 4 | 24.2 | 27.1 | 24 | 25.2 | 25.7 | 26.1 | 24.6 | 24.3 | 22 | 26.11 | 24.4121 | 26 | 26.5351 | 21.2 | 24.835 | 25 | 24.3 | 23 | 26.6 | 24 | 21.26 |  |  |  | 24.7736
| 1 | 5 | 24.4 | 26.7 | 23 | 24 | 25.3 | 25.3 | 24.5 | 25.7 | 22 | 25.67 | 23.4852 | 26.1 | 25.3547 | 21.7 | 25.167 | 25.3 | 22 | 23 | 26.1 | 24 | 22.2 |  |  |  | 102
| 1 | 6 | 24.5 | 26.7 | 24.2 | 23.7 | 26 | 25.7 | 25.7 | 25.7 | 21.9 | 25.44 | 23.8746 | 25.2 | 24.4651 | 21.9 | 25.602 | 25.3 | 23.8 | 22 | 26.3 | 24 | 21.14
| Mean |  | 24.15 | 26.5 | 24.0833 | 24.35 | 25.8167 | 25.5333 | 25.25 | 24.7 | 21.8667 | 25.78 | 23.9823 | 25.4833 | 25.5402 | 21.7167 | 24.9492 | 25.3667 | 23.7333 | 23 | 26.2667 | 24 | 21.29
| Median |  | 24.3 | 26.7 | 24.1 | 24.2 | 25.95 | 25.55 | 25.35 | 24.5 | 21.95 | 25.765 | 23.8925 | 25.4 | 25.4653 | 21.75 | 24.8595 | 25.3 | 23.95 | 23 | 26.3 | 24 | 21.32
| Std Dev. |  | 0.36742 | 0.73756 | 0.77567 | 0.55408 | 0.27869 | 0.37238 | 0.61563 | 0.86023 | 0.19664 | 0.28305 | 0.35118 | 0.50761 | 0.70078 | 0.27869 | 0.40277 | 0.3559 | 0.88015 | 0.63246 | 0.20656 | 0 | 0.60382
| Rel.Std.Dev. |  | 1.521% | 2.783% | 3.221% | 2.275% | 1.079% | 1.458% | 2.438% | 3.483% | 0.8993% | 1.098% | 1.464% | 1.992% | 2.744% | 1.283% | 1.614% | 1.403% | 3.709% | 2.75% | 0.7864% | 0% | 2.836%
| PDM3 |  | -2.517% | 6.969% | -2.786% | -1.71% | 4.21% | 3.067% | 1.923% | -0.2971% | -11.73% | 4.062% | -3.194% | 2.865% | 3.094% | -12.34% | 0.7087% | 2.394% | -4.199% | -7.159% | 6.027% | -3.123% | -14.06%
| Z-Score (Absolute) |  | 0.45 | 1.46 | 0.5 | 0.28 | 0.9 | 0.67 | 0.45 | 0 | 2.29 | 0.87 | 0.58 | 0.63 | 0.68 | 2.41 | 0.2 | 0.54 | 0.78 | 1.38 | 1.27 | 0.57 | 2.76
| Analytical results for Sr in OREAS 610 (Certified Value 38.6 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 1
| 1 | 1 | 49.2 | 34.6 | 51.4 | 37.8 | 38.6 | 37 | 37.6 | 38.5 | 49.4 | 40.05 | 54.1013 | 31 | 43 | 36.4 | 51.5909 | 29.8 | 37.54 | 39.1 | 38.9 | 38.1 | 33 | 29 | 61.86 |  | 1
| 1 | 2 | 49.2 | 36.1 | 50.1 | 37.2 | 39 | 38.8 | 38.7 | 39.4 | 49.4 | 39.61 | 55.964 | 30 | 40.2 | 36 | 51.0075 | 30.1 | 36.89 | 38.7 | 40.5 | 37.3 | 33 | 29 | 59.24 |  | 19
| 1 | 3 | 47.2 | 35.9 | 45.4 | 37.3 | 38 | 38.3 | 35.6 | 38.4 | 49.6 | 38.44 | 54.0135 | 30 | 42.2 | 33.7 | 53.831 | 30 | 36.73 | 39.1 | 38.9 | 39 | 31 | 29 | 56.62 |  | 16
| 1 | 4 | 50.6 | 34.8 | 48.2 | 37.7 | 36.6 | 39 | 35.8 | 39.4 | 49.9 | 38.18 | 56.2324 | 31 | 42.4 | 35.1 | 48.2791 | 28.5 | 35.74 | 38.7 | 39.7 | 38.3 | 31 | 29 | 54 |  | 38.6366
| 1 | 5 | 49.6 | 35.3 | 46 | 35.3 | 37.7 | 38.5 | 36.2 | 40.2 | 49.4 | 37.97 | 54.2384 | 31 | 43.4 | 33.2 | 48.3283 | 30.2 | 37.75 | 38.8 | 35.9 | 38.5 | 31 | 28 | 51.38 |  | 103
| 1 | 6 | 49.6 | 34.7 | 48.4 | 36.2 | 37.4 | 39.3 | 34.4 | 40.1 | 49.6 | 38.85 | 55.2659 | 31 | 41.9 | 31.7 | 46.459 | 30.3 | 40.53 | 38.1 | 38.6 | 38.7 | 30 | 29 | 48.76
| Mean |  | 49.2333 | 35.2333 | 48.25 | 36.9167 | 37.8833 | 38.4833 | 36.3833 | 39.3333 | 49.55 | 38.85 | 54.9693 | 30.6667 | 42.1833 | 34.35 | 49.916 | 29.8167 | 37.53 | 38.75 | 38.75 | 38.3167 | 31.5 | 28.8333 | 55.31
| Median |  | 49.4 | 35.05 | 48.3 | 37.25 | 37.85 | 38.65 | 36 | 39.4 | 49.5 | 38.645 | 54.7522 | 31 | 42.3 | 34.4 | 49.6679 | 30.05 | 37.215 | 38.75 | 38.9 | 38.4 | 31 | 29 | 55.31
| Std Dev. |  | 1.1201 | 0.64395 | 2.3046 | 0.97451 | 0.85888 | 0.8085 | 1.5342 | 0.76333 | 0.19748 | 0.82571 | 0.98732 | 0.5164 | 1.1143 | 1.8031 | 2.7009 | 0.66758 | 1.6311 | 0.36742 | 1.5592 | 0.58793 | 1.2247 | 0.40825 | 4.9016
| Rel.Std.Dev. |  | 2.275% | 1.828% | 4.776% | 2.64% | 2.267% | 2.101% | 4.217% | 1.941% | 0.3986% | 2.125% | 1.796% | 1.684% | 2.642% | 5.249% | 5.411% | 2.239% | 4.346% | 0.9482% | 4.024% | 1.534% | 3.888% | 1.416% | 8.862%
| PDM3 |  | 27.43% | -8.808% | 24.88% | -4.452% | -1.95% | -0.3968% | -5.832% | 1.803% | 28.25% | 0.5522% | 42.27% | -20.63% | 9.18% | -11.09% | 29.19% | -22.83% | -2.864% | 0.2934% | 0.2934% | -0.8281% | -18.47% | -25.37% | 43.15%
| Z-Score (Absolute) |  | 1.96 | 0.59 | 1.78 | 0.29 | 0.11 | 0 | 0.38 | 0.15 | 2.02 | 0.07 | 3 | 1.42 | 0.67 | 0.75 | 2.08 | 1.58 | 0.17 | 0.05 | 0.05 | 0.03 | 1.27 | 1.76 | 3.07
| Analytical results for Ta in OREAS 610 (Indicative Value < 0.01 ppm) | [25 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [7 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19477 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19363 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 9
| 1 | 3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19325 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19566 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19268 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 15
| 1 | 6 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19519 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1
| Mean |  | 2.12: 0.1942
| Median |  | 2.12: 0.1942
| Std Dev. |  | 2.12: 0.0011802
| Rel.Std.Dev. |  | 2.12: 0.6077%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 0 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0 | 0.67 | 5.7 | 0.67 | 167.73 | 4.21 | 0 | 0 | 0.84
| Analytical results for Tb in OREAS 610 (Indicative Value 0.17 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab | [7 empty] | 1
| No | No | 2.02: 2.04; 2.03: 2.18; 2.04: 2.20; 2.05: 2.21 | [7 empty] | ppm
|  |  | 2.02: AR*MS; 2.03: AR*OES/MS; 2.04: AR*OES/MS; 2.05: AR*OES/MS | [7 empty] | 2
|  |  | 2.02: 0.5g; 2.03: 0.2g; 2.04: 0.25g; 2.05: 0.25g | [7 empty] | 2
| 1 | 1 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.2; 2.05: 0.17 | [7 empty] | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.2; 2.05: 0.17 | [7 empty] | 10
| 1 | 3 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.19; 2.05: 0.17 | [7 empty] | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.19; 2.05: 0.16 | [7 empty] | 0.165
| 1 | 5 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.2; 2.05: 0.15 | [7 empty] | 16
| 1 | 6 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.2; 2.05: 0.16
| Mean |  | 2.02: 0.2; 2.03: 0.1; 2.04: 0.19667; 2.05: 0.16333
| Median |  | 2.02: 0.2; 2.03: 0.1; 2.04: 0.2; 2.05: 0.165
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 1.5202e-17; 2.04: 0.005164; 2.05: 0.008165
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 1.52e-14%; 2.04: 2.626%; 2.05: 4.999%
| PDM3 |  | 2.02: 21.21%; 2.03: -39.39%; 2.04: 19.19%; 2.05: -1.01%
| Z-Score (Absolute) |  | 2.02: 0.74; 2.03: 2.94; 2.04: 0.61; 2.05: 0.61
| Analytical results for Te in OREAS 610 (Certified Value 41.7 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | [5 empty] | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | [5 empty] | 2
| 1 | 1 | 39.73 | 43.93 | 40 | 43.1 | 45.8 | 39.8 | 44.4 | 42.5 | 38.7 | 42.04 | 44.437 | 38.53 | 43.7 | 43.7668 | 40.57 | 43.86 | 40.54 | 39.4 | 29 | [5 empty] | 1
| 1 | 2 | 38.65 | 43.93 | 38.3 | 41.4 | 45.7 | 41 | 43.7 | 40.9 | 39.2 | 40.99 | 43.8111 | 37.23 | 43.1 | 44.0096 | 41.2 | 43.39 | 39.07 | 38.5 | 30 | [5 empty] | 32
| 1 | 3 | 38.19 | 44.27 | 38 | 42.6 | 46.2 | 40.6 | 43.3 | 41.8 | 38.3 | 40.78 | 43.5657 | 39.13 | 43.3 | 43.1134 | 40.47 | 42.12 | 39.21 | 39.4 | 30 | [5 empty] | 16
| 1 | 4 | 39.49 | 44.78 | 37.9 | 42.1 | 45.5 | 42.2 | 42.4 | 42.3 | 38.7 | 41.34 | 45.3569 | 42.51 | 45.8 | 43.0936 | 40.6 | 42.63 | 38.92 | 39.6 | 29 | [5 empty] | 41.7054
| 1 | 5 | 39.09 | 44.41 | 36.6 | 42.1 | 44.7 | 40.4 | 42.2 | 44 | 38.8 | 41.36 | 43.0996 | 40.37 | 45.2 | 44.2137 | 41.9 | 43.03 | 37.02 | 39 | 29 | [5 empty] | 104
| 1 | 6 | 40.95 | 44.31 | 38.4 | 41.5 | 45.6 | 41.3 | 44.3 | 44.6 | 39 | 40.88 | 44.6593 | 40.16 | 44.2 | 41.7693 | 42.17 | 43.97 | 39.03 | 39.1 | 30
| Mean |  | 39.35 | 44.2717 | 38.2 | 42.1333 | 45.5833 | 40.8833 | 43.3833 | 42.6833 | 38.7833 | 41.2317 | 44.1549 | 39.655 | 44.2167 | 43.3277 | 41.1517 | 43.1667 | 38.965 | 39.1667 | 29.5
| Median |  | 39.29 | 44.29 | 38.15 | 42.1 | 45.65 | 40.8 | 43.5 | 42.4 | 38.75 | 41.165 | 44.1241 | 39.645 | 43.95 | 43.4401 | 40.9 | 43.21 | 39.05 | 39.25 | 29.5
| Std Dev. |  | 0.962 | 0.32015 | 1.0936 | 0.64704 | 0.49565 | 0.82563 | 0.93256 | 1.3819 | 0.30605 | 0.4621 | 0.81948 | 1.8077 | 1.0797 | 0.89132 | 0.73576 | 0.71807 | 1.1261 | 0.39328 | 0.54772
| Rel.Std.Dev. |  | 2.445% | 0.7232% | 2.863% | 1.536% | 1.087% | 2.019% | 2.15% | 3.238% | 0.7891% | 1.121% | 1.856% | 4.559% | 2.442% | 2.057% | 1.788% | 1.663% | 2.89% | 1.004% | 1.857%
| PDM3 |  | -5.648% | 6.153% | -8.405% | 1.026% | 9.298% | -1.971% | 4.023% | 2.345% | -7.006% | -1.136% | 5.874% | -4.916% | 6.022% | 3.89% | -1.328% | 3.504% | -6.571% | -6.087% | -29.27%
| Z-Score (Absolute) |  | 0.72 | 0.8 | 1.08 | 0.14 | 1.21 | 0.25 | 0.53 | 0.31 | 0.9 | 0.14 | 0.76 | 0.63 | 0.78 | 0.51 | 0.16 | 0.46 | 0.84 | 0.78 | NA
Results from laboratory 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 610 (Certified Value 3.08 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.25 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 01g |  |  |  |  | 3
| 1 | 1 | 4.67 | 3 | 3.5 | 2.8 | 2.8 | 2.8 | 2.9 | 3.1 | 4.93 | 3.372 | 4.1461 | 6 | 3.5 | 3 | < 10 | 2.87 | 3.3 | 3.3 | 3.2 | < 2 |  |  |  |  | 1
| 1 | 2 | 4.73 | 3.1 | 3.4 | 3 | 2.9 | 2.9 | 3 | 3.1 | 4.92 | 3.291 | 4.1959 | 5 | 3.1 | 3.2 | < 10 | 2.79 | 3.3 | 3.4 | 3.2 | < 2 |  |  |  |  | 33
| 1 | 3 | 4.69 | 3.1 | 3.4 | 2.9 | 2.8 | 2.9 | 2.9 | 3.2 | 4.94 | 3.507 | 4.2347 | 6 | 3.1 | 3.2 | < 10 | 2.73 | 3.4 | 3.3 | 3.1 | < 2 |  |  |  |  | 16
| 1 | 4 | 4.8 | 3.1 | 3.4 | 3 | 2.7 | 3 | 2.9 | 3.2 | 4.61 | 3.295 | 4.212 | 6 | 3.5 | 3.1 | < 10 | 2.68 | 3.3 | 3.2 | 2.9 | < 2 |  |  |  |  | 3.0845
| 1 | 5 | 4.86 | 3 | 3.2 | 2.8 | 2.8 | 2.9 | 2.8 | 3.3 | 4.8 | 3.372 | 4.1591 | 6 | 3.4 | 3 | < 10 | 2.76 | 3.3 | 3 | 2.9 | < 2 |  |  |  |  | 105
| 1 | 6 | 4.79 | 3 | 3.4 | 2.9 | 2.7 | 3 | 3 | 3.2 | 5.04 | 3.275 | 4.2042 | 6 | 3.3 | 3 | < 10 | 2.66 | 3.2 | 3.2 | 2.8 | < 2
| Mean |  | 4.7567 | 3.05 | 3.3833 | 2.9 | 2.7833 | 2.9167 | 2.9167 | 3.1833 | 4.8733 | 3.352 | 4.192 | 5.8333 | 3.3167 | 3.0833 |  | 2.7483 | 3.3 | 3.2333 | 3.0167
| Median |  | 4.76 | 3.05 | 3.4 | 2.9 | 2.8 | 2.9 | 2.9 | 3.2 | 4.925 | 3.3335 | 4.2 | 6 | 3.35 | 3.05 |  | 2.745 | 3.3 | 3.25 | 3
| Std Dev. |  | 0.072572 | 0.054772 | 0.098319 | 0.089443 | 0.075277 | 0.075277 | 0.075277 | 0.075277 | 0.14989 | 0.086861 | 0.033407 | 0.40825 | 0.18348 | 0.098319 |  | 0.07679 | 0.063246 | 0.13663 | 0.17224
| Rel.Std.Dev. |  | 1.526% | 1.796% | 2.906% | 3.084% | 2.705% | 2.581% | 2.581% | 2.365% | 3.076% | 2.591% | 0.7969% | 6.999% | 5.532% | 3.189% |  | 2.794% | 1.917% | 4.226% | 5.71%
| PDM3 |  | 54.21% | -1.12% | 9.687% | -5.983% | -9.765% | -5.443% | -5.443% | 3.203% | 57.99% | 8.671% | 35.9% | 89.11% | 7.525% | -0.03937% |  | -10.9% | 6.985% | 4.824% | -2.201%
| Z-Score (Absolute) |  | 3.98 | 0.34 | 0.51 | 0.72 | 1.01 | 0.67 | 0.67 | 0 | 4.27 | 0.43 | 2.55 | NA | 0.34 | 0.25 | 4.59 | 1.1 | 0.3 | 0.13 | 0.42 | 5.52
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 610 (Indicative Value 0.008 wt.%) | [25 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.0124 | 0.007 | 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0149 | 0.0079 | 0.015487 | 0.009 | 0.008 | 0.006 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.007 | 0.01 | 0.02651 |  | 1
| 1 | 2 | 0.0125 | 0.007 | 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.015 | 0.0077 | 0.016029 | 0.005 | 0.008 | 0.006 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.007 | 0.01 | 0.025 |  | 11
| 1 | 3 | 0.0122 | 0.007 | 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0146 | 0.007 | 0.015824 | 0.006 | 0.008 | 0.006 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.007 | 0.01 | 0.02649 |  | 16
| 1 | 4 | 0.013 | 0.007 | 0.01 | 0.008 | 0.008 | 0.007 | 0.008 | 0.008 | 0.0143 | 0.007 | 0.015827 | 0.006 | 0.008 | 0.006 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.006 | 0.01 | 0.02598 |  | 0.0081878
| 1 | 5 | 0.0126 | 0.007 | 0.01 | 0.008 | 0.008 | 0.008 | 0.007 | 0.008 | 0.0152 | 0.0068 | 0.015618 | 0.006 | 0.009 | 0.006 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.006 | 0.01 | 0.02647 |  | 17
| 1 | 6 | 0.0126 | 0.007 | 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.0154 | 0.0074 | 0.015944 | 0.006 | 0.008 | 0.005 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | < 0.01 | 0.006 | 0.01 | 0.02696
| Mean |  | 0.01255 | 0.007 | 0.01 | 0.008 | 0.008 | 0.0078333 | 0.0078333 | 0.008 | 0.0149 | 0.0073 | 0.015788 | 0.0063333 | 0.0081667 | 0.0058333 |  |  |  | 0.01 |  |  | 0.0065 | 0.01 | 0.026235
| Median |  | 0.01255 | 0.007 | 0.01 | 0.008 | 0.008 | 0.008 | 0.008 | 0.008 | 0.01495 | 0.0072 | 0.015825 | 0.006 | 0.008 | 0.006 |  |  |  | 0.01 |  |  | 0.0065 | 0.01 | 0.02648
| Std Dev. |  | 0.00026646 | 9.5015e-19 | 0 | 0 | 0 | 0.00040825 | 0.00040825 | 0 | 0.0004 | 0.00043818 | 0.00020249 | 0.0013663 | 0.00040825 | 0.00040825 |  |  |  | 0 |  |  | 0.00054772 | 0 | 0.00067996
| Rel.Std.Dev. |  | 2.123% | 1.357e-14% | 0% | 0% | 0% | 5.212% | 5.212% | 0% | 2.685% | 6.002% | 1.283% | 21.57% | 4.999% | 6.999% |  |  |  | 0% |  |  | 8.427% | 0% | 2.592%
| PDM3 |  | 53.28% | -14.51% | 22.13% | -2.293% | -2.293% | -4.329% | -4.329% | -2.293% | 81.98% | -10.84% | 92.82% | -22.65% | -0.2578% | -28.76% |  |  |  | 22.13% |  |  | -20.61% | 22.13% | 220.4%
| Z-Score (Absolute) |  | 1.47 | 0.26 | 0.67 | 0.05 | 0.05 | 0 | 0 | 0.05 | 2.2 | 0.17 | 2.48 | 0.47 | 0.1 | 0.62 | 0.88 | 0.88 | 0.88 | 0.67 | 0.88 | 0.88 | 0.41 | 0.67 | 5.73
| Analytical results for Tl in OREAS 610 (Certified Value 1.49 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.25 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 01g |  |  |  |  | 3
| 1 | 1 | 1.51 | 1.51 | 1.59 | 1.45 | 1.48 | 1.43 | 1.42 | 1.48 | 1.51 | 1.5 | 1.5317 | < 5 | 1.46 | 1.18 | < 5 | 1.19 | 1.48 | 1.45 | 1.7 | < 5 |  |  |  |  | 1
| 1 | 2 | 1.46 | 1.56 | 1.55 | 1.43 | 1.47 | 1.43 | 1.39 | 1.46 | 1.53 | 1.45 | 1.577 | < 5 | 1.43 | 1.19 | < 5 | 1.19 | 1.55 | 1.45 | 1.7 | < 5 |  |  |  |  | 35
| 1 | 3 | 1.46 | 1.55 | 1.53 | 1.44 | 1.48 | 1.45 | 1.41 | 1.5 | 1.5 | 1.52 | 1.5206 | < 5 | 1.39 | 1.21 | < 5 | 1.18 | 1.5 | 1.45 | 1.7 | < 5 |  |  |  |  | 16
| 1 | 4 | 1.51 | 1.57 | 1.56 | 1.54 | 1.43 | 1.5 | 1.38 | 1.53 | 1.52 | 1.48 | 1.5521 | < 5 | 1.52 | 1.16 | < 5 | 1.17 | 1.51 | 1.41 | 1.7 | < 5 |  |  |  |  | 1.4893
| 1 | 5 | 1.53 | 1.55 | 1.47 | 1.46 | 1.45 | 1.44 | 1.39 | 1.56 | 1.52 | 1.49 | 1.539 | < 5 | 1.44 | 1.14 | < 5 | 1.19 | 1.53 | 1.29 | 1.7 | < 5 |  |  |  |  | 106
| 1 | 6 | 1.56 | 1.55 | 1.54 | 1.38 | 1.52 | 1.47 | 1.47 | 1.51 | 1.56 | 1.46 | 1.5613 | < 5 | 1.47 | 1.22 | < 5 | 1.18 | 1.56 | 1.43 | 1.7 | < 5
| Mean |  | 1.505 | 1.5483 | 1.54 | 1.45 | 1.4717 | 1.4533 | 1.41 | 1.5067 | 1.5233 | 1.4833 | 1.5469 |  | 1.4517 | 1.1833 |  | 1.1833 | 1.5217 | 1.4133 | 1.7
| Median |  | 1.51 | 1.55 | 1.545 | 1.445 | 1.475 | 1.445 | 1.4 | 1.505 | 1.52 | 1.485 | 1.5455 |  | 1.45 | 1.185 |  | 1.185 | 1.52 | 1.44 | 1.7
| Std Dev. |  | 0.03937 | 0.020412 | 0.04 | 0.052154 | 0.030605 | 0.027325 | 0.032863 | 0.03559 | 0.020656 | 0.02582 | 0.020615 |  | 0.043551 | 0.030111 |  | 0.008165 | 0.030605 | 0.062503 | 0
| Rel.Std.Dev. |  | 2.616% | 1.318% | 2.597% | 3.597% | 2.08% | 1.88% | 2.331% | 2.362% | 1.356% | 1.741% | 1.333% |  | 3% | 2.545% |  | 0.69% | 2.011% | 4.422% | 0%
| PDM3 |  | 1.055% | 3.965% | 3.406% | -2.638% | -1.183% | -2.414% | -5.323% | 1.167% | 2.286% | -0.3994% | 3.872% |  | -2.526% | -20.54% |  | -20.54% | 2.175% | -5.1% | 14.15%
| Z-Score (Absolute) |  | 0.01 | 0.54 | 0.43 | 0.71 | 0.43 | 0.66 | 1.21 | 0.01 | 0.22 | 0.28 | 0.52 | 12.57 | 0.68 | 4.08 | 12.57 | 4.08 | 0.2 | 1.17 | 2.45 | 12.57
| Analytical results for Tm in OREAS 610 (Indicative Value < 100 ppb) | [25 empty] | Indicative
Tm | Round | Replicate | 2.02: Lab |  |  |  |  | 1
| No | No | 2.02: 2.04 |  |  |  |  | ppb
|  |  | 2.02: AR*MS |  |  |  |  | 0
|  |  | 2.02: 0.5g |  |  |  |  | 0
| 1 | 1 | 2.02: < 100 |  |  |  |  | 1
| 1 | 2 | 2.02: < 100 |  |  |  |  | 6
| 1 | 3 | 2.02: < 100 |  |  |  |  | 16
| 1 | 4 | 2.02: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.02: < 100 |  |  |  |  | 18
| 1 | 6 | 2.02: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.02
| Analytical results for U in OREAS 610 (Certified Value 1.12 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [6 empty] | 3
| 1 | 1 | 1.19 | 1.2 | 1.3 | 1.04 | 1.04 | 1.04 | 1.09 | 1.15 | 1.22 | 1.155 | 1.1543 | 1.07 | 0.99 | < 10 | 0.83 | 1.11 | 1.13 | 1.2 | [6 empty] | 1
| 1 | 2 | 1.16 | 1.2 | 1.2 | 1.04 | 1.07 | 1.06 | 1.09 | 1.11 | 1.22 | 1.122 | 1.1914 | 1 | 0.98 | < 10 | 0.84 | 1.14 | 1.17 | 1.2 | [6 empty] | 37
| 1 | 3 | 1.21 | 1.2 | 1.2 | 0.98 | 1.03 | 1.06 | 1.09 | 1.15 | 1.22 | 1.184 | 1.1978 | 1 | 0.94 | < 10 | 0.81 | 1.13 | 1.13 | 1.2 | [6 empty] | 16
| 1 | 4 | 1.2 | 1.2 | 1.2 | 1.05 | 1.03 | 1.12 | 1.09 | 1.18 | 1.2 | 1.142 | 1.1728 | 1.1 | 0.97 | < 10 | 0.81 | 1.13 | 1.1 | 1.1 | [6 empty] | 1.1203
| 1 | 5 | 1.22 | 1.2 | 1.2 | 1.02 | 1.05 | 1.08 | 1.05 | 1.19 | 1.18 | 1.147 | 1.1621 | 1.11 | 0.95 | < 10 | 0.82 | 1.14 | 1.02 | 1.1 | [6 empty] | 107
| 1 | 6 | 1.21 | 1.2 | 1.2 | 1 | 1.04 | 1.05 | 1.04 | 1.18 | 1.18 | 1.135 | 1.2016 | 1.04 | 0.98 | < 10 | 0.83 | 1.15 | 1.1 | 1.1
| Mean |  | 1.1983 | 1.2 | 1.2167 | 1.0217 | 1.0433 | 1.0683 | 1.075 | 1.16 | 1.2033 | 1.1475 | 1.18 | 1.0533 | 0.96833 |  | 0.82333 | 1.1333 | 1.1083 | 1.15
| Median |  | 1.205 | 1.2 | 1.2 | 1.03 | 1.04 | 1.06 | 1.09 | 1.165 | 1.21 | 1.1445 | 1.1821 | 1.055 | 0.975 |  | 0.825 | 1.135 | 1.115 | 1.15
| Std Dev. |  | 0.02137 | 0 | 0.040825 | 0.027142 | 0.015055 | 0.028577 | 0.023452 | 0.029665 | 0.019664 | 0.021097 | 0.019717 | 0.048028 | 0.019408 |  | 0.012111 | 0.013663 | 0.050365 | 0.054772
| Rel.Std.Dev. |  | 1.783% | 0% | 3.355% | 2.657% | 1.443% | 2.675% | 2.182% | 2.557% | 1.634% | 1.839% | 1.671% | 4.56% | 2.004% |  | 1.471% | 1.206% | 4.544% | 4.763%
| PDM3 |  | 6.969% | 7.118% | 8.606% | -8.801% | -6.867% | -4.635% | -4.04% | 3.547% | 7.416% | 2.432% | 5.332% | -5.974% | -13.56% |  | -26.51% | 1.167% | -1.065% | 2.655%
| Z-Score (Absolute) |  | 0.61 | 0.63 | 0.8 | 1.25 | 1.02 | 0.76 | 0.69 | 0.21 | 0.66 | 0.07 | 0.42 | 0.92 | 1.81 | 40.56 | 3.33 | 0.07 | 0.34 | 0.1
| Analytical results for V in OREAS 610 (Certified Value 11.6 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 2
| 1 | 1 | 14 | 12 | 13 | 11 | 11 | 11 | 11 | 11 | 14 | 12 | 9.4784 | 10 | 12 | 12.6 | 11.5333 | 9.7 | 11.48 | 13 | 12 | 13 | 12 | 12 | 14.61 |  | 1
| 1 | 2 | 13 | 11 | 13 | 11 | 12 | 11 | 11 | 11 | 14 | 11.9 | 9.7225 | 10 | 12 | 13.1 | 11.4379 | 9.7 | 11.464 | 13 | 11 | 13 | 12 | 12 | 13.61 |  | 38
| 1 | 3 | 14 | 12 | 12 | 11 | 11 | 11 | 11 | 11 | 14 | 11.1 | 9.3807 | 10 | 11 | 12.8 | 11.6366 | 9.3 | 11.344 | 13 | 11 | 13 | 12 | 12 | 13.81 |  | 16
| 1 | 4 | 13 | 11 | 13 | 11 | 11 | 11 | 11 | 11 | 14 | 11 | 9.7071 | 10 | 13 | 12.9 | 11.163 | 9.3 | 11.319 | 13 | 11 | 13 | 12 | 11 | 13.41 |  | 11.5915
| 1 | 5 | 14 | 11 | 12 | 11 | 12 | 11 | 11 | 12 | 14 | 11 | 9.4168 | 10 | 11 | 12.7 | 11.2067 | 9.3 | 11.393 | 13 | 11 | 13 | 11 | 11 | 13.01 |  | 108
| 1 | 6 | 14 | 12 | 13 | 11 | 11 | 11 | 11 | 11 | 14 | 11.7 | 9.3154 | 10 | 11 | 12.6 | 11.0325 | 9.3 | 11.468 | 12 | 11 | 13 | 12 | 11 | 13.11
| Mean |  | 13.6667 | 11.5 | 12.6667 | 11 | 11.3333 | 11 | 11 | 11.1667 | 14 | 11.45 | 9.5035 | 10 | 11.6667 | 12.7833 | 11.335 | 9.4333 | 11.4113 | 12.8333 | 11.1667 | 13 | 11.8333 | 11.5 | 13.5933
| Median |  | 14 | 11.5 | 13 | 11 | 11 | 11 | 11 | 11 | 14 | 11.4 | 9.4476 | 10 | 11.5 | 12.75 | 11.3223 | 9.3 | 11.4285 | 13 | 11 | 13 | 12 | 11.5 | 13.51
| Std Dev. |  | 0.5164 | 0.54772 | 0.5164 | 0 | 0.5164 | 0 | 0 | 0.40825 | 0 | 0.46797 | 0.17207 | 0 | 0.8165 | 0.19408 | 0.23597 | 0.20656 | 0.069419 | 0.40825 | 0.40825 | 0 | 0.40825 | 0.54772 | 0.58109
| Rel.Std.Dev. |  | 3.779% | 4.763% | 4.077% | 0% | 4.556% | 0% | 0% | 3.656% | 0% | 4.087% | 1.811% | 0% | 6.999% | 1.518% | 2.082% | 2.19% | 0.6083% | 3.181% | 3.656% | 0% | 3.45% | 4.763% | 4.275%
| PDM3 |  | 17.9% | -0.7891% | 9.276% | -5.103% | -2.227% | -5.103% | -5.103% | -3.665% | 20.78% | -1.22% | -18.01% | -13.73% | 0.6488% | 10.28% | -2.212% | -18.62% | -1.554% | 10.71% | -3.665% | 12.15% | 2.087% | -0.7891% | 17.27%
| Z-Score (Absolute) |  | 3.32 | 0.07 | 1.82 | 0.67 | 0.17 | 0.67 | 0.67 | 0.42 | 3.82 | 0 | 2.92 | 2.17 | 0.32 | 2 | 0.17 | 3.02 | 0.06 | 2.07 | 0.42 | 2.32 | 0.57 | 0.07 | 3.21
| Analytical results for W in OREAS 610 (Certified Value 3.58 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 01g |  |  | 2
| 1 | 1 | 2.9 | 3.4 | 4.3 | 3.5 | 3.83 | 3.38 | 3.67 | 4.05 | 2.8 | 3.57 | 3.8648 | < 2 | 5.87 | 4.26 | < 20 | 3.4 | < 10 | 3.8 | 3.3 | < 10 | 3.5 | < 5 |  |  | 1
| 1 | 2 | 2.8 | 3.4 | 4.2 | 3.75 | 4.01 | 3.36 | 3.68 | 3.88 | 2.7 | 3.56 | 3.5399 | < 2 | 5.62 | 4.25 | < 20 | 3.4 | < 10 | 4 | 3.3 | < 10 | 3.5 | < 5 |  |  | 39
| 1 | 3 | 2.9 | 3.3 | 3.8 | 3.61 | 3.9 | 3.28 | 3.62 | 3.96 | 2.8 | 3.59 | 3.8572 | < 2 | 5.34 | 4.36 | < 20 | 3.3 | < 10 | 4 | 3.3 | < 10 | 3.6 | < 5 |  |  | 16
| 1 | 4 | 3 | 3.3 | 4.2 | 3.71 | 3.84 | 3.46 | 3.63 | 4.01 | 2.7 | 3.59 | 3.8502 | < 2 | 6.05 | 4.19 | < 20 | 3.3 | < 10 | 3.8 | 3.2 | < 10 | 3.5 | < 5 |  |  | 3.5799
| 1 | 5 | 2.9 | 3.2 | 4.3 | 3.6 | 3.96 | 3.39 | 3.56 | 4.13 | 2.7 | 3.56 | 3.5202 | < 2 | 5.9 | 4.16 | < 20 | 3.4 | < 10 | 4 | 3 | < 10 | 3.5 | < 5 |  |  | 109
| 1 | 6 | 2.9 | 3.4 | 4.1 | 3.62 | 3.83 | 3.4 | 4.02 | 4.19 | 2.8 | 3.47 | 3.6452 | < 2 | 5.78 | 4.27 | < 20 | 3.3 | < 10 | 3.9 | 3.1 | < 10 | 3.5 | < 5
| Mean |  | 2.9 | 3.3333 | 4.15 | 3.6317 | 3.895 | 3.3783 | 3.6967 | 4.0367 | 2.75 | 3.5567 | 3.7129 |  | 5.76 | 4.2483 |  | 3.35 |  | 3.9167 | 3.2 |  | 3.5167
| Median |  | 2.9 | 3.35 | 4.2 | 3.615 | 3.87 | 3.385 | 3.65 | 4.03 | 2.75 | 3.565 | 3.7477 |  | 5.825 | 4.255 |  | 3.35 |  | 3.95 | 3.25 |  | 3.5
| Std Dev. |  | 0.063246 | 0.08165 | 0.18708 | 0.088412 | 0.076092 | 0.058793 | 0.16403 | 0.11272 | 0.054772 | 0.044572 | 0.16395 |  | 0.24988 | 0.06969 |  | 0.054772 |  | 0.098319 | 0.12649 |  | 0.040825
| Rel.Std.Dev. |  | 2.181% | 2.449% | 4.508% | 2.434% | 1.954% | 1.74% | 4.437% | 2.792% | 1.992% | 1.253% | 4.416% |  | 4.338% | 1.64% |  | 1.635% |  | 2.51% | 3.953% |  | 1.161%
| PDM3 |  | -18.99% | -6.887% | 15.93% | 1.446% | 8.802% | -5.63% | 3.262% | 12.76% | -23.18% | -0.6487% | 3.716% |  | 60.9% | 18.67% |  | -6.422% |  | 9.407% | -10.61% |  | -1.766%
| Z-Score (Absolute) |  | 1.23 | 0.53 | 0.78 | 0.05 | 0.37 | 0.46 | 0.05 | 0.6 | 1.47 | 0.17 | 0.08 | 4.29 | 3.38 | 0.94 | 10.21 | 0.51 | 2.15 | 0.41 | 0.75 | 2.15 | 0.24 | 1.88
| Analytical results for Y in OREAS 610 (Certified Value 3.09 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g |  |  |  |  | 3
| 1 | 1 | 3.4 | 2.93 | 3.39 | 3.17 | 3.13 | 2.93 | 3.16 | 3.14 | 3.25 | 3.17 | 3.092 | 2.85 | 2.9016 | 3.1 | 3.43 | 3.17 | 3 | < 5 | 3.6 | 3 |  |  |  |  | 1
| 1 | 2 | 3.34 | 3.19 | 3.32 | 2.95 | 3.12 | 3.06 | 3.19 | 3.11 | 3.22 | 3.13 | 3.1505 | 2.8 | 2.8712 | 3.1 | 3.37 | 3.11 | 2.97 | < 5 | 3.6 | 3 |  |  |  |  | 40
| 1 | 3 | 3.22 | 3.01 | 3.09 | 3.04 | 3.11 | 2.99 | 3.03 | 3.08 | 3.33 | 3.05 | 3.0529 | 2.77 | 2.9008 | 3.1 | 3.39 | 3.07 | 2.96 | < 5 | 3.6 | 3 |  |  |  |  | 16
| 1 | 4 | 3.4 | 3.06 | 3.24 | 2.95 | 3.06 | 3.15 | 3.02 | 3.17 | 3.19 | 3.07 | 3.0715 | 2.89 | 2.7965 | 3.1 | 3.42 | 3.03 | 2.95 | < 5 | 3.5 | 3 |  |  |  |  | 3.0932
| 1 | 5 | 3.41 | 3.11 | 3.11 | 2.94 | 2.98 | 3.02 | 3.11 | 3.26 | 3.26 | 3.04 | 2.9651 | 2.85 | 2.8145 | 3.1 | 3.41 | 3.06 | 2.66 | < 5 | 3.5 | 3 |  |  |  |  | 110
| 1 | 6 | 3.41 | 3.07 | 3.32 | 2.93 | 3.11 | 3.05 | 2.81 | 3.26 | 3.14 | 3.04 | 3.0332 | 2.71 | 2.7731 | 3.1 | 3.43 | 3.07 | 2.91 | < 5 | 3.4 | 3
| Mean |  | 3.3633 | 3.0617 | 3.245 | 2.9967 | 3.085 | 3.0333 | 3.0533 | 3.17 | 3.2317 | 3.0833 | 3.0609 | 2.8117 | 2.8429 | 3.1 | 3.4083 | 3.085 | 2.9083 |  | 3.5333 | 3
| Median |  | 3.4 | 3.065 | 3.28 | 2.95 | 3.11 | 3.035 | 3.07 | 3.155 | 3.235 | 3.06 | 3.0622 | 2.825 | 2.8429 | 3.1 | 3.415 | 3.07 | 2.955 |  | 3.55 | 3
| Std Dev. |  | 0.075011 | 0.088185 | 0.12211 | 0.093737 | 0.056833 | 0.073937 | 0.13721 | 0.075895 | 0.064936 | 0.054283 | 0.061823 | 0.065243 | 0.055541 | 0 | 0.024014 | 0.048888 | 0.12513 |  | 0.08165 | 0
| Rel.Std.Dev. |  | 2.23% | 2.88% | 3.763% | 3.128% | 1.842% | 2.437% | 4.494% | 2.394% | 2.009% | 1.761% | 2.02% | 2.32% | 1.954% | 0% | 0.7046% | 1.585% | 4.302% |  | 2.311% | 0%
| PDM3 |  | 8.734% | -1.019% | 4.908% | -3.12% | -0.2646% | -1.935% | -1.288% | 2.483% | 4.477% | -0.3185% | -1.045% | -9.101% | -8.09% | 0.2203% | 10.19% | -0.2646% | -5.976% |  | 14.23% | -3.013%
| Z-Score (Absolute) |  | 2.18 | 0.17 | 1.26 | 0.67 | 0.01 | 0.39 | 0.23 | 0.67 | 1.15 | 0 | 0.17 | 2.11 | 1.87 | 0.13 | 2.53 | 0.01 | 1.36 | 4.54 | 3.5 | NA
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 610 (Indicative Value 0.2 ppm) | [25 empty] | Indicative
Yb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab |  |  |  |  | 1
| No | No | 2.02: 2.04; 2.03: 2.18; 2.04: 2.20; 2.05: 2.21 |  |  |  |  | ppm
|  |  | 2.02: AR*MS; 2.03: AR*OES/MS; 2.04: AR*OES/MS; 2.05: AR*OES/MS |  |  |  |  | 2
|  |  | 2.02: 0.5g; 2.03: 0.2g; 2.04: 0.25g; 2.05: 0.25g |  |  |  |  | 2
| 1 | 1 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 13
| 1 | 3 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 0.2
| 1 | 5 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 19
| 1 | 6 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2
| Mean |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2
| Median |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 3.0405e-17; 2.04: 3.0405e-17; 2.05: 3.0405e-17
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 1.52e-14%; 2.04: 1.52e-14%; 2.05: 1.52e-14%
| PDM3 |  | 2.02: -1.11e-14%; 2.03: -1.11e-14%; 2.04: -1.11e-14%; 2.05: -1.11e-14%
Z-Score (Absolute) | bar: 2.02 to 2.05
| Analytical results for Zn in OREAS 610 (Certified Value 1764 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 0
| 1 | 1 | 2060 | 1872 | 1848.5 | 1780 | 1760 | 1780 | 1710 | 1770 | 1740 | 1770 | 1834.1 | 1667.9004 | 1817 | 1749.6 | 1850 | 1830.0371 | 1738 | 1793.805 | 1745 | 1858 | 1730 | 1918 | 1760 | 1609 | 1
| 1 | 2 | 1960 | 1774 | 1865.4 | 1740 | 1760 | 1790 | 1700 | 1770 | 1660 | 1800 | 1820.1 | 1687.9394 | 1817 | 1710.9 | 1840 | 1819.5262 | 1736 | 1778.852 | 1728 | 1924 | 1750 | 1935 | 1750 | 1596 | 22
| 1 | 3 | 1950 | 1837 | 1864.3 | 1650 | 1750 | 1770 | 1690 | 1800 | 1720 | 1790 | 1739.8 | 1600.5433 | 1787 | 1736.6 | 1780 | 1836.3123 | 1748 | 1787.95 | 1681 | 1868 | 1770 | 1951 | 1750 | 1603 | 16
| 1 | 4 | 2080 | 1797 | 1863 | 1710 | 1730 | 1780 | 1700 | 1760 | 1720 | 1760 | 1745.9 | 1681.631 | 1834 | 1749.5 | 1830 | 1838.1728 | 1700 | 1793.549 | 1703 | 1875 | 1730 | 2000 | 1700 | 1570 | 1763.5635
| 1 | 5 | 1910 | 1828 | 1865.1 | 1610 | 1750 | 1780 | 1690 | 1720 | 1800 | 1760 | 1755.1 | 1617.9342 | 1836 | 1753.2 | 1750 | 1831.0814 | 1742 | 1785.776 | 1694 | 1782 | 1730 | 1950 | 1710 | 1587 | 111
| 1 | 6 | 2050 | 1838 | 1861.1 | 1670 | 1720 | 1760 | 1700 | 1720 | 1740 | 1740 | 1756 | 1639.8651 | 1835 | 1725.5 | 1730 | 1830.5645 | 1728 | 1772.867 | 1700 | 1899 | 1750 | 1960 | 1720 | 1584
| Mean |  | 2001.6667 | 1824.3333 | 1861.2333 | 1693.3333 | 1745 | 1776.6667 | 1698.3333 | 1756.6667 | 1730 | 1770 | 1775.1667 | 1649.3022 | 1821 | 1737.55 | 1796.6667 | 1830.949 | 1732 | 1785.4665 | 1708.5 | 1867.6667 | 1743.3333 | 1952.3333 | 1731.6667 | 1591.5
| Median |  | 2005 | 1832.5 | 1863.65 | 1690 | 1750 | 1780 | 1700 | 1765 | 1730 | 1765 | 1755.55 | 1653.8827 | 1825.5 | 1743.05 | 1805 | 1830.8229 | 1737 | 1786.863 | 1701.5 | 1871.5 | 1740 | 1950.5 | 1735 | 1591.5
| Std Dev. |  | 70.2614 | 34.3899 | 6.4342 | 62.1825 | 16.4317 | 10.328 | 7.5277 | 31.4113 | 45.1664 | 21.9089 | 40.9134 | 35.593 | 18.8574 | 16.6689 | 50.4645 | 6.5181 | 17.0176 | 8.29 | 23.586 | 48.2521 | 16.3299 | 27.6309 | 24.8328 | 14.1244
| Rel.Std.Dev. |  | 3.51% | 1.885% | 0.3457% | 3.672% | 0.9416% | 0.5813% | 0.4432% | 1.788% | 2.611% | 1.238% | 2.305% | 2.158% | 1.036% | 0.9593% | 2.809% | 0.356% | 0.9825% | 0.4643% | 1.381% | 2.584% | 0.9367% | 1.415% | 1.434% | 0.8875%
| PDM3 |  | 13.5% | 3.446% | 5.538% | -3.982% | -1.053% | 0.743% | -3.699% | -0.3911% | -1.903% | 0.365% | 0.6579% | -6.479% | 3.257% | -1.475% | 1.877% | 3.821% | -1.79% | 1.242% | -3.122% | 5.903% | -1.147% | 10.7% | -1.809% | -9.757%
| Z-Score (Absolute) |  | 3.65 | 0.93 | 1.5 | 1.07 | 0.28 | 0.2 | 0.99 | 0.1 | 0.51 | 0.1 | 0.18 | 1.74 | 0.88 | 0.39 | 0.51 | 1.03 | 0.48 | 0.34 | 0.84 | 1.6 | 0.31 | 2.89 | 0.48 | 2.63
| Analytical results for Zr in OREAS 610 (Certified Value 11.1 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  |  | 2
| 1 | 1 | 12.2 | 11 | 14.7 | 11.9 | 11.8 | 11.5 | 12 | 11.7 | 10.6 | 12.71 | 9.4719 | 10.7 | 12.1005 | 9.8 | 10.48 | 10.4 | 10.5 | 10.7 | 20.2 | 11 | 11.63 |  |  |  | 1
| 1 | 2 | 12 | 11.2 | 14.3 | 10.8 | 11.9 | 11.9 | 11.9 | 11.6 | 10.4 | 12.65 | 10.006 | 10.8 | 11.9421 | 9.7 | 10.49 | 10.1 | 10.8 | 10.9 | 20.1 | 11 | 12.4 |  |  |  | 23
| 1 | 3 | 11.8 | 10.9 | 14.9 | 11.1 | 11.9 | 11.7 | 11.5 | 11.8 | 10.8 | 12.38 | 9.3817 | 10.3 | 12.2188 | 9.9 | 10.71 | 9.9 | 10.7 | 10.9 | 19.9 | 11 | 12.17 |  |  |  | 16
| 1 | 4 | 12.4 | 11.1 | 14.2 | 11.4 | 11.5 | 12.2 | 11.5 | 12 | 10.3 | 12.47 | 9.9823 | 10.9 | 11.4028 | 9.8 | 10.56 | 10.2 | 10.5 | 10.5 | 19.2 | 10 | 10.24 |  |  |  | 11.0958
| 1 | 5 | 12.3 | 11.1 | 13.5 | 11 | 11.4 | 11.7 | 11.5 | 12.4 | 10.4 | 12.31 | 9.5618 | 10.5 | 11.5537 | 10 | 10.89 | 10.1 | 9.4 | 11.1 | 18.8 | 10 | 10.71 |  |  |  | 112
| 1 | 6 | 12.4 | 11 | 15.1 | 11 | 11.7 | 11.8 | 10.2 | 12.4 | 10.7 | 12.4 | 9.5949 | 10.2 | 11.0973 | 9.8 | 10.95 | 10.2 | 10.3 | 10.7 | 18.6 | 10 | 10.48
| Mean |  | 12.1833 | 11.05 | 14.45 | 11.2 | 11.7 | 11.8 | 11.4333 | 11.9833 | 10.5333 | 12.4867 | 9.6664 | 10.5667 | 11.7192 | 9.8333 | 10.68 | 10.15 | 10.3667 | 10.8 | 19.4667 | 10.5 | 11.2717
| Median |  | 12.25 | 11.05 | 14.5 | 11.05 | 11.75 | 11.75 | 11.5 | 11.9 | 10.5 | 12.435 | 9.5783 | 10.6 | 11.7479 | 9.8 | 10.635 | 10.15 | 10.5 | 10.8 | 19.55 | 10.5 | 11.17
| Std Dev. |  | 0.24014 | 0.10488 | 0.57879 | 0.39497 | 0.20976 | 0.23664 | 0.64395 | 0.34881 | 0.19664 | 0.15933 | 0.2646 | 0.28048 | 0.43793 | 0.10328 | 0.20416 | 0.16432 | 0.50465 | 0.20976 | 0.69186 | 0.54772 | 0.91816
| Rel.Std.Dev. |  | 1.971% | 0.9491% | 4.005% | 3.527% | 1.793% | 2.005% | 5.632% | 2.911% | 1.867% | 1.276% | 2.737% | 2.654% | 3.737% | 1.05% | 1.912% | 1.619% | 4.868% | 1.942% | 3.554% | 5.216% | 8.146%
| PDM3 |  | 9.802% | -0.4125% | 30.23% | 0.9393% | 5.446% | 6.347% | 3.042% | 7.999% | -5.069% | 12.54% | -12.88% | -4.769% | 5.618% | -11.38% | -3.747% | -8.524% | -6.571% | -2.666% | 75.44% | -5.369% | 1.585%
| Z-Score (Absolute) |  | 0.93 | 0.18 | 3.16 | 0.04 | 0.46 | 0.55 | 0.19 | 0.73 | 0.69 | 1.23 | 1.54 | 0.66 | 0.48 | 1.38 | 0.55 | 1.07 | 0.85 | 0.43 | 8.09 | NA | 0.04
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C513: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C550: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C568: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C605: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C642: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C660: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C805: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C823: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C841: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C897: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C969: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1134: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 610 (Indicative Value 11.72 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.7 | 1
| 1 | 2 | 11.74 | 16
| Mean |  | 11.72 | 16
| Median |  | 11.72 | 11.72
| Std Dev. |  | 0.028284 | 22
| Rel.Std.Dev. |  | 0.2413%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 610 (Indicative Value 3035 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 3030 | 1
| 1 | 2 | 3040 | 17
| Mean |  | 3035 | 16
| Median |  | 3035 | 3035
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 0.233%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 610 (Indicative Value 4615 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 4630 | 1
| 1 | 2 | 4600 | 18
| Mean |  | 4615 | 16
| Median |  | 4615 | 4615
| Std Dev. |  | 21.2132 | 24
| Rel.Std.Dev. |  | 0.4597%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 610 (Indicative Value 0.33 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.33 | 1
| 1 | 2 | 0.33 | 19
| Mean |  | 0.33 | 16
| Median |  | 0.33 | 0.33
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 610 (Indicative Value 35 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 40 | 20
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 7.0711 | 26
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 610 (Indicative Value 15 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 20 | 21
| Mean |  | 15 | 16
| Median |  | 15 | 15
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 610 (Indicative Value 65 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 70 | 22
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 7.0711 | 28
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 610 (Indicative Value 0.969 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.965 | 1
| 1 | 2 | 0.972 | 23
| Mean |  | 0.9685 | 16
| Median |  | 0.9685 | 0.9685
| Std Dev. |  | 0.0049497 | 29
| Rel.Std.Dev. |  | 0.5111%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 610 (Indicative Value 3.39 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.38 | 1
| 1 | 2 | 3.39 | 16
| Mean |  | 3.385 | 16
| Median |  | 3.385 | 3.385
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 0.2089%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 610 (Indicative Value 2.41 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.4 | 1
| 1 | 2 | 2.41 | 17
| Mean |  | 2.405 | 16
| Median |  | 2.405 | 2.405
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 0.294%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 610 (Indicative Value 0.29 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.29 | 1
| 1 | 2 | 0.29 | 18
| Mean |  | 0.29 | 16
| Median |  | 0.29 | 0.29
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 610 (Indicative Value 0.011 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.01 | 1
| 1 | 2 | 0.011 | 19
| Mean |  | 0.0105 | 16
| Median |  | 0.0105 | 0.0105
| Std Dev. |  | 0.00070711 | 25
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 610 (Indicative Value 1.15 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.14 | 1
| 1 | 2 | 1.16 | 20
| Mean |  | 1.15 | 16
| Median |  | 1.15 | 1.15
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 1.23%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 610 (Indicative Value 35 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 40 | 21
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 610 (Indicative Value 0.125 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.126 | 1
| 1 | 2 | 0.124 | 22
| Mean |  | 0.125 | 16
| Median |  | 0.125 | 0.125
| Std Dev. |  | 0.0014142 | 28
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 610 (Indicative Value 710 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 700 | 1
| 1 | 2 | 720 | 23
| Mean |  | 710 | 16
| Median |  | 710 | 710
| Std Dev. |  | 14.1421 | 29
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 610 (Indicative Value 4.2 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.2047 | 1
| 1 | 2 | 4.2047 | 16
| Mean |  | 4.2047 | 16
| Median |  | 4.2047 | 4.2047
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -0.0002206%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 610 (Indicative Value 69.21 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69.18 | 1
| 1 | 2 | 69.24 | 17
| Mean |  | 69.21 | 16
| Median |  | 69.21 | 69.21
| Std Dev. |  | 0.042426 | 23
| Rel.Std.Dev. |  | 0.0613%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 610 (Indicative Value 10 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 18
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 610 (Indicative Value 338 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 340 | 1
| 1 | 2 | 340 | 19
| Mean |  | 340 | 16
| Median |  | 340 | 338.2378
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.521%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 610 (Indicative Value 0.278 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.276 | 1
| 1 | 2 | 0.28 | 20
| Mean |  | 0.278 | 16
| Median |  | 0.278 | 0.278
| Std Dev. |  | 0.0028284 | 26
| Rel.Std.Dev. |  | 1.017%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 610 (Indicative Value 65 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 70 | 21
| Mean |  | 65 | 16
| Median |  | 65 | 65
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 610 (Indicative Value 1795 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 1790 | 1
| 1 | 2 | 1800 | 22
| Mean |  | 1795 | 16
| Median |  | 1795 | 1795
| Std Dev. |  | 7.0711 | 28
| Rel.Std.Dev. |  | 0.3939%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 610 (Indicative Value 144 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 141 | 1
| 1 | 2 | 148 | 23
| Mean |  | 144.5 | 16
| Median |  | 144.5 | 144.3628
| Std Dev. |  | 4.9497 | 29
| Rel.Std.Dev. |  | 3.425%
| PDM3 |  | 0.09505%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 610 (Indicative Value 8.74 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 8.75 | 1
| 1 | 2 | 8.73 | 25
| Mean |  | 8.74 | 16
| Median |  | 8.74 | 8.74
| Std Dev. |  | 0.014142 | 31
| Rel.Std.Dev. |  | 0.1618%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 610 (Indicative Value 57.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 56.6 | 1
| 1 | 2 | 57.8 | 27
| Mean |  | 57.2 | 16
| Median |  | 57.2 | 57.2
| Std Dev. |  | 0.84853 | 33
| Rel.Std.Dev. |  | 1.483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 610 (Indicative Value 2980 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2950 | 1
| 1 | 2 | 3010 | 17
| Mean |  | 2980 | 16
| Median |  | 2980 | 2980
| Std Dev. |  | 42.4264 | 34
| Rel.Std.Dev. |  | 1.424%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 610 (Indicative Value 3865 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 3820 | 1
| 1 | 2 | 3910 | 5
| Mean |  | 3865 | 16
| Median |  | 3865 | 3865
| Std Dev. |  | 63.6396 | 35
| Rel.Std.Dev. |  | 1.647%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 610 (Indicative Value 1.8 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.8 | 1
| 1 | 2 | 1.8 | 30
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 610 (Indicative Value 220 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 219 | 1
| 1 | 2 | 221 | 31
| Mean |  | 220 | 16
| Median |  | 220 | 220
| Std Dev. |  | 1.4142 | 37
| Rel.Std.Dev. |  | 0.6428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 610 (Indicative Value 14 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13.5 | 1
| 1 | 2 | 14.5 | 32
| Mean |  | 14 | 16
| Median |  | 14 | 14
| Std Dev. |  | 0.70711 | 38
| Rel.Std.Dev. |  | 5.051%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 610 (Indicative Value 49.5 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 49.3 | 1
| 1 | 2 | 49.6 | 33
| Mean |  | 49.45 | 16
| Median |  | 49.45 | 49.45
| Std Dev. |  | 0.21213 | 39
| Rel.Std.Dev. |  | 0.429%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 610 (Indicative Value 7.45 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.6 | 1
| 1 | 2 | 7.3 | 21
| Mean |  | 7.45 | 16
| Median |  | 7.45 | 7.45
| Std Dev. |  | 0.21213 | 40
| Rel.Std.Dev. |  | 2.847%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 610 (Indicative Value 46.5 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 45 | 1
| 1 | 2 | 48 | 35
| Mean |  | 46.5 | 16
| Median |  | 46.5 | 46.5
| Std Dev. |  | 2.1213 | 41
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 610 (Indicative Value 2.12 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2.13 | 36
| Mean |  | 2.115 | 16
| Median |  | 2.115 | 2.115
| Std Dev. |  | 0.021213 | 42
| Rel.Std.Dev. |  | 1.003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 610 (Indicative Value 0.973 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.955 | 1
| 1 | 2 | 0.991 | 23
| Mean |  | 0.973 | 16
| Median |  | 0.973 | 0.973
| Std Dev. |  | 0.025456 | 43
| Rel.Std.Dev. |  | 2.616%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 610 (Indicative Value 1.54 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.55 | 10
| Mean |  | 1.54 | 16
| Median |  | 1.54 | 1.54
| Std Dev. |  | 0.014142 | 44
| Rel.Std.Dev. |  | 0.9183%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 610 (Indicative Value 0.6 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.59 | 1
| 1 | 2 | 0.61 | 11
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0.014142 | 45
| Rel.Std.Dev. |  | 2.357%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 610 (Indicative Value 0.78 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.78 | 1
| 1 | 2 | 0.78 | 12
| Mean |  | 0.78 | 16
| Median |  | 0.78 | 0.78
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 610 (Indicative Value 23.1 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23 | 1
| 1 | 2 | 23.2 | 41
| Mean |  | 23.1 | 16
| Median |  | 23.1 | 23.1
| Std Dev. |  | 0.14142 | 47
| Rel.Std.Dev. |  | 0.6122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 610 (Indicative Value 2.79 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.79 | 1
| 1 | 2 | 2.78 | 13
| Mean |  | 2.785 | 16
| Median |  | 2.785 | 2.785
| Std Dev. |  | 0.0070711 | 48
| Rel.Std.Dev. |  | 0.2539%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 610 (Indicative Value 4.55 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 4.5 | 6
| Mean |  | 4.55 | 16
| Median |  | 4.55 | 4.55
| Std Dev. |  | 0.070711 | 49
| Rel.Std.Dev. |  | 1.554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 610 (Indicative Value 3.82 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.82 | 1
| 1 | 2 | 3.82 | 27
| Mean |  | 3.82 | 16
| Median |  | 3.82 | 3.82
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 610 (Indicative Value 0.25 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.25 | 8
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 610 (Indicative Value 3.53 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.5 | 1
| 1 | 2 | 3.55 | 29
| Mean |  | 3.525 | 16
| Median |  | 3.525 | 3.525
| Std Dev. |  | 0.035355 | 35
| Rel.Std.Dev. |  | 1.003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 610 (Indicative Value 26.2 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 25.9 | 1
| 1 | 2 | 26.4 | 30
| Mean |  | 26.15 | 16
| Median |  | 26.15 | 26.15
| Std Dev. |  | 0.35355 | 36
| Rel.Std.Dev. |  | 1.352%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 610 (Indicative Value 85 ppb) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 90 | 1
| 1 | 2 | 80 | 6
| Mean |  | 85 | 16
| Median |  | 85 | 85
| Std Dev. |  | 7.0711 | 37
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 610 (Indicative Value 73 ppm) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72 | 1
| 1 | 2 | 73 | 32
| Mean |  | 72.5 | 16
| Median |  | 72.5 | 72.5
| Std Dev. |  | 0.70711 | 38
| Rel.Std.Dev. |  | 0.9753%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 610 (Indicative Value 4.6 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 4.6 | 33
| Mean |  | 4.6 | 16
| Median |  | 4.6 | 4.6
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 610 (Indicative Value 8.37 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.31 | 1
| 1 | 2 | 8.42 | 34
| Mean |  | 8.365 | 16
| Median |  | 8.365 | 8.365
| Std Dev. |  | 0.077782 | 40
| Rel.Std.Dev. |  | 0.9298%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 610 (Indicative Value 20.1 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.9 | 1
| 1 | 2 | 20.2 | 10
| Mean |  | 20.05 | 16
| Median |  | 20.05 | 20.05
| Std Dev. |  | 0.21213 | 41
| Rel.Std.Dev. |  | 1.058%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 610 (Indicative Value 29 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 28 | 1
| 1 | 2 | 30 | 21
| Mean |  | 29 | 16
| Median |  | 29 | 29
| Std Dev. |  | 1.4142 | 42
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 610 (Indicative Value 732 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 724 | 1
| 1 | 2 | 739 | 23
| Mean |  | 731.5 | 16
| Median |  | 731.5 | 731.5
| Std Dev. |  | 10.6066 | 43
| Rel.Std.Dev. |  | 1.45%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 610 (Indicative Value 5.45 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.46 | 1
| 1 | 2 | 5.43 | 12
| Mean |  | 5.445 | 16
| Median |  | 5.445 | 5.445
| Std Dev. |  | 0.021213 | 44
| Rel.Std.Dev. |  | 0.3896%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 610 (Indicative Value 62 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 61.5 | 1
| 1 | 2 | 62.5 | 39
| Mean |  | 62 | 16
| Median |  | 62 | 62
| Std Dev. |  | 0.70711 | 45
| Rel.Std.Dev. |  | 1.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 610 (Indicative Value < 10 ppb) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | < 10 | 1
| 1 | 2 | < 10 | 14
| Mean |  |  | 16
| Median |  |  | < 10
| Std Dev. |  |  | 46
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 610 (Indicative Value 312 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 309 | 1
| 1 | 2 | 315 | 41
| Mean |  | 312 | 16
| Median |  | 312 | 312
| Std Dev. |  | 4.2426 | 47
| Rel.Std.Dev. |  | 1.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 610 (Indicative Value 3.55 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.6 | 1
| 1 | 2 | 3.5 | 42
| Mean |  | 3.55 | 16
| Median |  | 3.55 | 3.55
| Std Dev. |  | 0.070711 | 48
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 610 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 43
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 49
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 610 (Indicative Value 3.94 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.94 | 1
| 1 | 2 | 3.94 | 8
| Mean |  | 3.94 | 16
| Median |  | 3.94 | 3.94
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 610 (Indicative Value 26.4 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.4 | 1
| 1 | 2 | 26.4 | 18
| Mean |  | 26.4 | 16
| Median |  | 26.4 | 26.4
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 610 (Indicative Value 333 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 332 | 1
| 1 | 2 | 334 | 19
| Mean |  | 333 | 16
| Median |  | 333 | 333
| Std Dev. |  | 1.4142 | 35
| Rel.Std.Dev. |  | 0.4247%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 610 (Indicative Value 0.73 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.75 | 1
| 1 | 2 | 0.71 | 9
| Mean |  | 0.73 | 16
| Median |  | 0.73 | 0.73
| Std Dev. |  | 0.028284 | 36
| Rel.Std.Dev. |  | 3.875%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 610 (Indicative Value 0.36 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.36 | 10
| Mean |  | 0.36 | 16
| Median |  | 0.36 | 0.36
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 610 (Indicative Value 45.3 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 44 | 1
| 1 | 2 | 46.6 | 32
| Mean |  | 45.3 | 16
| Median |  | 45.3 | 45.3
| Std Dev. |  | 1.8385 | 38
| Rel.Std.Dev. |  | 4.058%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 610 (Indicative Value 10.1 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10.1 | 33
| Mean |  | 10.05 | 16
| Median |  | 10.05 | 10.05
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.7036%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 610 (Indicative Value 0.164 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.163 | 1
| 1 | 2 | 0.165 | 11
| Mean |  | 0.164 | 16
| Median |  | 0.164 | 0.164
| Std Dev. |  | 0.0014142 | 40
| Rel.Std.Dev. |  | 0.8623%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 610 (Indicative Value 2.7 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.6 | 35
| Mean |  | 2.7 | 16
| Median |  | 2.7 | 2.7
| Std Dev. |  | 0.14142 | 41
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 610 (Indicative Value 105 ppb) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 110 | 6
| Mean |  | 105 | 16
| Median |  | 105 | 105
| Std Dev. |  | 7.0711 | 42
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 610 (Indicative Value 2.63 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.61 | 1
| 1 | 2 | 2.65 | 37
| Mean |  | 2.63 | 16
| Median |  | 2.63 | 2.63
| Std Dev. |  | 0.028284 | 43
| Rel.Std.Dev. |  | 1.075%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 610 (Indicative Value 31 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.2 | 1
| 1 | 2 | 30.7 | 38
| Mean |  | 30.95 | 16
| Median |  | 30.95 | 30.95
| Std Dev. |  | 0.35355 | 44
| Rel.Std.Dev. |  | 1.142%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 610 (Indicative Value 7.25 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.5 | 1
| 1 | 2 | 7 | 39
| Mean |  | 7.25 | 16
| Median |  | 7.25 | 7.25
| Std Dev. |  | 0.35355 | 45
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 610 (Indicative Value 7.11 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.08 | 1
| 1 | 2 | 7.14 | 40
| Mean |  | 7.11 | 16
| Median |  | 7.11 | 7.11
| Std Dev. |  | 0.042426 | 46
| Rel.Std.Dev. |  | 0.5967%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 610 (Indicative Value 0.65 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.65 | 1
| 1 | 2 | 0.65 | 13
| Mean |  | 0.65 | 16
| Median |  | 0.65 | 0.65
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 610 (Indicative Value 1765 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1780 | 1
| 1 | 2 | 1750 | 22
| Mean |  | 1765 | 16
| Median |  | 1765 | 1765
| Std Dev. |  | 21.2132 | 48
| Rel.Std.Dev. |  | 1.202%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 610 (Indicative Value 126 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 125 | 1
| 1 | 2 | 126 | 23
| Mean |  | 125.5 | 16
| Median |  | 125.5 | 125.5
| Std Dev. |  | 0.70711 | 49
| Rel.Std.Dev. |  | 0.5634%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## XRPA
| Analytical results for Au in OREAS 610 (Indicative Value 9.87 ppm) |  |  |  | Indicative
Au | Round | Replicate | Chrysos | Chrysos | 1
| No | No | Perth | Kalgoorlie | ppm
|  |  | Raw*XRPA | Raw*XRPA | 2
2
| 1 | 1 | 9.938 | 9.795 | 1
| Mean |  | 9.938 | 9.795 | 45
| Median |  | 9.938 | 9.795 | 16
| Std Dev. |  |  |  | 9.8665
| Rel.Std.Dev. |  |  |  | 51
| PDM3 |  | 0.7247% | -0.7247%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C12: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.