## Abbreviations
Table 1. Abbreviations used for OREAS 611
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with AAS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with either AAS or ICP-OES finish as appropriate
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
ABL*MS | Laser Ablation with Mass Spectrometry: ICP-MS finish
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/AAS | ore grade aqua regia digest with either AAS or ICP-OES finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
FA*AAS | fire assay with AAS finish
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
LOI*TGA | loss on ignition with Thermal Gravimetric Analyser finish
PF*OES | sodium peroxide fusion with ICP-OES finish
Raw*XRPA | raw sample (solid) packed into single use jar by X-Ray Photon Assay
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 611
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
CERTIMIN, Lima, Peru
Chrysos Corporation Limited, Kalgoorlie, WA, Australia
Chrysos Corporation Limited, Perth, WA, Australia
Inspectorate (BV), Lima, Peru
Inspectorate America Corporation (BV), Sparks, Nevada, USA
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
On Site Laboratory Services, Bendigo, VIC, Australia
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
Skyline Assayers & Laboratories, Tucson, Arizona, USA

## Performance Gates
Table 1. Pooled-Lab Performance Gates for OREAS 611
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 15.6989 | 0.60102 | 14.4969 | 16.9009 | 13.8959 | 17.502 | 3.828% | 7.657% | 11.49% | 14.914 | 16.4839
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 15.5317 | 0.40669 | 14.7183 | 16.3451 | 14.3116 | 16.7518 | 2.618% | 5.237% | 7.855% | 14.7551 | 16.3083
Infrared Combustion
S, wt.% | 4.2615 | 0.17658 | 3.9083 | 4.6146 | 3.7317 | 4.7912 | 4.144% | 8.287% | 12.43% | 4.0484 | 4.4745
4-Acid Digestion
Ag, ppm | 80.0064 | 1.613 | 76.7804 | 83.2325 | 75.1673 | 84.8456 | 2.016% | 4.032% | 6.048% | 76.0061 | 84.0068
Al, wt.% | 5.7836 | 0.21424 | 5.3551 | 6.2121 | 5.1408 | 6.4263 | 3.704% | 7.409% | 11.11% | 5.4944 | 6.0727
As, ppm | 3399.6771 | 179.6305 | 3040.4162 | 3758.938 | 2860.7858 | 3938.5685 | 5.284% | 10.57% | 15.85% | 3229.6933 | 3569.661
Be, ppm | 1.4613 | 0.1064 | 1.2485 | 1.6741 | 1.1421 | 1.7805 | 7.281% | 14.56% | 21.84% | 1.3882 | 1.5344
Bi, ppm | 265.4204 | 10.9916 | 243.4372 | 287.4036 | 232.4456 | 298.3952 | 4.141% | 8.282% | 12.42% | 252.1494 | 278.6914
Ca, wt.% | 0.22729 | 0.01767 | 0.19195 | 0.26263 | 0.17428 | 0.2803 | 7.774% | 15.55% | 23.32% | 0.21592 | 0.23865
Cd, ppm | 13.912 | 0.73495 | 12.4421 | 15.3819 | 11.7072 | 16.1169 | 5.283% | 10.57% | 15.85% | 13.2164 | 14.6076
Ce, ppm | 46.4045 | 4.4604 | 37.4837 | 55.3254 | 33.0233 | 59.7858 | 9.612% | 19.22% | 28.84% | 44.0843 | 48.7247
Co, ppm | 8.6962 | 0.26628 | 8.1636 | 9.2288 | 7.8974 | 9.495 | 3.062% | 6.124% | 9.186% | 8.2614 | 9.131
Cr, ppm | 50.5179 | 7.1553 | 36.2074 | 64.8285 | 29.0521 | 71.9838 | 14.16% | 28.33% | 42.49% | 47.992 | 53.0438
Cs, ppm | 2.0619 | 0.14766 | 1.7665 | 2.3572 | 1.6189 | 2.5048 | 7.162% | 14.32% | 21.48% | 1.9588 | 2.1649
Cu, wt.% | 1.1667 | 0.021569 | 1.1236 | 1.2099 | 1.102 | 1.2314 | 1.849% | 3.697% | 5.546% | 1.1084 | 1.2251
Dy, ppm | 1.5513 | 0.1096 | 1.3321 | 1.7705 | 1.2225 | 1.8801 | 7.065% | 14.13% | 21.19% | 1.4738 | 1.6289
Er, ppm | 0.613 | 0.041514 | 0.52997 | 0.69603 | 0.48846 | 0.73754 | 6.772% | 13.54% | 20.32% | 0.58235 | 0.64365
Eu, ppm | 0.75567 | 0.054989 | 0.64569 | 0.86565 | 0.5907 | 0.92063 | 7.277% | 14.55% | 21.83% | 0.71788 | 0.79345
Fe, wt.% | 2.5366 | 0.092493 | 2.3517 | 2.7216 | 2.2592 | 2.8141 | 3.646% | 7.293% | 10.94% | 2.4098 | 2.6635
Ga, ppm | 24.439 | 1.6781 | 21.0827 | 27.7952 | 19.4046 | 29.4733 | 6.867% | 13.73% | 20.6% | 23.217 | 25.6609
Gd, ppm | 2.9173 | 0.2339 | 2.4495 | 3.3851 | 2.2156 | 3.619 | 8.018% | 16.04% | 24.05% | 2.7715 | 3.0632
Hf, ppm | 2.2641 | 0.11733 | 2.0295 | 2.4988 | 1.9121 | 2.6161 | 5.182% | 10.36% | 15.55% | 2.1509 | 2.3773
Ho, ppm | 0.21867 | 0.023887 | 0.17089 | 0.26644 | 0.14701 | 0.29033 | 10.92% | 21.85% | 32.77% | 0.20773 | 0.2296
In, ppm | 4.6787 | 0.26784 | 4.1431 | 5.2144 | 3.8752 | 5.4823 | 5.725% | 11.45% | 17.17% | 4.4448 | 4.9127
K, wt.% | 1.8594 | 0.069803 | 1.7198 | 1.999 | 1.65 | 2.0688 | 3.754% | 7.508% | 11.26% | 1.7664 | 1.9523
La, ppm | 20.7995 | 3.6259 | 13.5478 | 28.0513 | 9.9219 | 31.6771 | 17.43% | 34.86% | 52.3% | 19.7596 | 21.8395
Li, ppm | 29.6931 | 1.5518 | 26.5895 | 32.7966 | 25.0377 | 34.3484 | 5.226% | 10.45% | 15.68% | 28.2084 | 31.1777
Mg, ppm | 1472.4424 | 93.1452 | 1286.152 | 1658.7328 | 1193.0068 | 1751.878 | 6.326% | 12.65% | 18.98% | 1398.8202 | 1546.0645
Mn, ppm | 78.8582 | 3.7586 | 71.341 | 86.3754 | 67.5824 | 90.134 | 4.766% | 9.533% | 14.3% | 74.9153 | 82.8011
Mo, ppm | 5.4576 | 0.37019 | 4.7172 | 6.198 | 4.347 | 6.5681 | 6.783% | 13.57% | 20.35% | 5.1847 | 5.7304
Na, wt.% | 0.80378 | 0.028317 | 0.74714 | 0.86041 | 0.71882 | 0.88873 | 3.523% | 7.046% | 10.57% | 0.76359 | 0.84397
Nb, ppm | 8.8228 | 0.56945 | 7.6839 | 9.9617 | 7.1144 | 10.5311 | 6.454% | 12.91% | 19.36% | 8.3817 | 9.2639
Nd, ppm | 19.393 | 1.2368 | 16.9193 | 21.8667 | 15.6825 | 23.1035 | 6.378% | 12.76% | 19.13% | 18.4233 | 20.3626
Ni, ppm | 27.0308 | 1.0901 | 24.8505 | 29.211 | 23.7604 | 30.3011 | 4.033% | 8.066% | 12.1% | 25.6792 | 28.3823
P, ppm | 547.8208 | 22.2632 | 503.2944 | 592.3472 | 481.0312 | 614.6104 | 4.064% | 8.128% | 12.19% | 520.4298 | 575.2118
Pb, ppm | 639.5673 | 34.8959 | 569.7754 | 709.3592 | 534.8795 | 744.2551 | 5.456% | 10.91% | 16.37% | 607.5889 | 671.5457
Pr, ppm | 5.1463 | 0.58641 | 3.9735 | 6.3192 | 3.3871 | 6.9056 | 11.39% | 22.79% | 34.18% | 4.889 | 5.4036
Rb, ppm | 64.5088 | 2.9125 | 58.6839 | 70.3337 | 55.7714 | 73.2462 | 4.515% | 9.03% | 13.54% | 61.2834 | 67.7343
S, wt.% | 4.177 | 0.16284 | 3.8513 | 4.5027 | 3.6885 | 4.6655 | 3.898% | 7.797% | 11.7% | 3.9681 | 4.3858
Sb, ppm | 364.7518 | 30.63 | 303.4917 | 426.0119 | 272.8617 | 456.6419 | 8.398% | 16.8% | 25.19% | 346.5142 | 382.9894
Sc, ppm | 3.2 | 0.22418 | 2.7516 | 3.6483 | 2.5274 | 3.8725 | 7.006% | 14.01% | 21.02% | 3.04 | 3.36
Se, ppm | 34.7806 | 3.4058 | 27.969 | 41.5921 | 24.5633 | 44.9979 | 9.792% | 19.58% | 29.38% | 33.0415 | 36.5196
Sm, ppm | 3.8013 | 0.22495 | 3.3514 | 4.2512 | 3.1265 | 4.4762 | 5.918% | 11.84% | 17.75% | 3.6113 | 3.9914
Sn, ppm | 32.0502 | 2.3449 | 27.3605 | 36.7399 | 25.0157 | 39.0848 | 7.316% | 14.63% | 21.95% | 30.4477 | 33.6527
Sr, ppm | 316.8633 | 21.3289 | 274.2056 | 359.5211 | 252.8767 | 380.8499 | 6.731% | 13.46% | 20.19% | 301.0202 | 332.7065
Ta, ppm | 0.71394 | 0.04496 | 0.62402 | 0.80386 | 0.57906 | 0.84882 | 6.297% | 12.59% | 18.89% | 0.67825 | 0.74964
Tb, ppm | 0.31079 | 0.065442 | 0.17991 | 0.44167 | 0.11447 | 0.50712 | 21.06% | 42.11% | 63.17% | 0.29525 | 0.32633
Te, ppm | 49.2816 | 2.7822 | 43.7171 | 54.8461 | 40.9349 | 57.6284 | 5.646% | 11.29% | 16.94% | 46.8176 | 51.7457
Th, ppm | 8.6627 | 1.0794 | 6.504 | 10.8214 | 5.4246 | 11.9008 | 12.46% | 24.92% | 37.38% | 8.2296 | 9.0958
Ti, wt.% | 0.19187 | 0.006893 | 0.17809 | 0.20566 | 0.17119 | 0.21255 | 3.593% | 7.185% | 10.78% | 0.18228 | 0.20147
Tl, ppm | 2.1746 | 0.11418 | 1.9463 | 2.403 | 1.8321 | 2.5172 | 5.251% | 10.5% | 15.75% | 2.0659 | 2.2833
U, ppm | 2.6771 | 0.16446 | 2.3482 | 3.006 | 2.1837 | 3.1704 | 6.143% | 12.29% | 18.43% | 2.5432 | 2.8109
V, ppm | 31.8691 | 1.498 | 28.8732 | 34.865 | 27.3752 | 36.363 | 4.7% | 9.401% | 14.1% | 30.2757 | 33.4626
W, ppm | 8.7494 | 0.45404 | 7.8413 | 9.6574 | 7.3872 | 10.1115 | 5.189% | 10.38% | 15.57% | 8.3119 | 9.1868
Y, ppm | 6.8039 | 0.33754 | 6.1288 | 7.479 | 5.7913 | 7.8165 | 4.961% | 9.922% | 14.88% | 6.4637 | 7.1441
Yb, ppm | 0.55042 | 0.077238 | 0.39594 | 0.70489 | 0.3187 | 0.78213 | 14.03% | 28.07% | 42.1% | 0.5229 | 0.57794
Zn, ppm | 2022.9122 | 72.8504 | 1877.2115 | 2168.6129 | 1804.3612 | 2241.4633 | 3.601% | 7.203% | 10.8% | 1921.7666 | 2124.0579
Zr, ppm | 68.8064 | 5.3153 | 58.1758 | 79.437 | 52.8605 | 84.7523 | 7.725% | 15.45% | 23.17% | 65.3661 | 72.2467
Aqua Regia Digestion
Ag, ppm | 79.2496 | 3.6206 | 72.0084 | 86.4907 | 68.3878 | 90.1113 | 4.569% | 9.137% | 13.71% | 75.2871 | 83.212
Al, wt.% | 0.83827 | 0.04819 | 0.74189 | 0.93465 | 0.6937 | 0.98284 | 5.749% | 11.5% | 17.25% | 0.79636 | 0.88019
As, ppm | 3340.4787 | 192.0963 | 2956.2861 | 3724.6713 | 2764.1898 | 3916.7676 | 5.751% | 11.5% | 17.25% | 3173.4548 | 3507.5026
Be, ppm | 0.28063 | 0.024551 | 0.23152 | 0.32973 | 0.20697 | 0.35428 | 8.749% | 17.5% | 26.25% | 0.26659 | 0.29466
Bi, ppm | 256.488 | 15.897 | 224.6941 | 288.2819 | 208.7971 | 304.1789 | 6.198% | 12.4% | 18.59% | 243.6636 | 269.3124
Ca, wt.% | 0.10798 | 0.0048076 | 0.098365 | 0.1176 | 0.093558 | 0.1224 | 4.452% | 8.905% | 13.36% | 0.10258 | 0.11338
Cd, ppm | 13.7293 | 0.97814 | 11.7731 | 15.6856 | 10.7949 | 16.6638 | 7.124% | 14.25% | 21.37% | 13.0429 | 14.4158
Ce, ppm | 14.6378 | 1.45 | 11.7378 | 17.5377 | 10.2878 | 18.9877 | 9.906% | 19.81% | 29.72% | 13.9059 | 15.3696
Co, ppm | 8.6376 | 0.46892 | 7.6998 | 9.5755 | 7.2308 | 10.0444 | 5.429% | 10.86% | 16.29% | 8.2057 | 9.0695
Cr, ppm | 41.6612 | 2.9096 | 35.842 | 47.4804 | 32.9324 | 50.39 | 6.984% | 13.97% | 20.95% | 39.5781 | 43.7442
Cs, ppm | 0.73252 | 0.07014 | 0.59224 | 0.8728 | 0.5221 | 0.94294 | 9.575% | 19.15% | 28.73% | 0.6959 | 0.76915
Cu, wt.% | 1.1791 | 0.023749 | 1.1316 | 1.2266 | 1.1078 | 1.2503 | 2.014% | 4.028% | 6.043% | 1.1201 | 1.238
Fe, wt.% | 2.4047 | 0.1238 | 2.1571 | 2.6523 | 2.0333 | 2.7761 | 5.148% | 10.3% | 15.44% | 2.2844 | 2.5249
Ga, ppm | 6.83 | 0.50387 | 5.8223 | 7.8378 | 5.3184 | 8.3417 | 7.377% | 14.75% | 22.13% | 6.4885 | 7.1715
Hf, ppm | 0.41875 | 0.030734 | 0.35728 | 0.48022 | 0.32655 | 0.51095 | 7.34% | 14.68% | 22.02% | 0.39781 | 0.43969
Hg, ppm | 0.94658 | 0.045683 | 0.85522 | 1.038 | 0.80953 | 1.0836 | 4.826% | 9.652% | 14.48% | 0.89926 | 0.99391
In, ppm | 4.5275 | 0.17201 | 4.1834 | 4.8715 | 4.0114 | 5.0435 | 3.799% | 7.598% | 11.4% | 4.3011 | 4.7538
K, wt.% | 0.2007 | 0.01152 | 0.17766 | 0.22374 | 0.16614 | 0.23526 | 5.74% | 11.48% | 17.22% | 0.19067 | 0.21074
La, ppm | 6.8578 | 0.66678 | 5.5242 | 8.1913 | 4.8574 | 8.8581 | 9.723% | 19.45% | 29.17% | 6.5149 | 7.2007
Li, ppm | 8.2534 | 0.61915 | 7.0151 | 9.4917 | 6.3959 | 10.1108 | 7.502% | 15% | 22.51% | 7.8407 | 8.666
Mg, ppm | 999.8813 | 67.3214 | 865.2385 | 1134.5242 | 797.9171 | 1201.8456 | 6.733% | 13.47% | 20.2% | 949.8873 | 1049.8754
Mn, ppm | 61.9307 | 3.5685 | 54.7936 | 69.0678 | 51.2251 | 72.6364 | 5.762% | 11.52% | 17.29% | 58.8342 | 65.0273
Mo, ppm | 5.1262 | 0.3195 | 4.4872 | 5.7652 | 4.1677 | 6.0847 | 6.233% | 12.47% | 18.7% | 4.8699 | 5.3825
Na, wt.% | 0.044336 | 0.0081792 | 0.027978 | 0.060695 | 0.019799 | 0.068874 | 18.45% | 36.9% | 55.34% | 0.042119 | 0.046553
Nb, ppm | 0.16821 | 0.030118 | 0.10797 | 0.22845 | 0.077853 | 0.25856 | 17.91% | 35.81% | 53.72% | 0.1598 | 0.17662
Ni, ppm | 26.9125 | 1.7578 | 23.3969 | 30.428 | 21.6391 | 32.1858 | 6.531% | 13.06% | 19.59% | 25.5668 | 28.2581
P, ppm | 242.618 | 12.1583 | 218.3014 | 266.9345 | 206.1431 | 279.0928 | 5.011% | 10.02% | 15.03% | 230.4871 | 254.7489
Pb, ppm | 480.6527 | 24.6983 | 431.2561 | 530.0493 | 406.5578 | 554.7476 | 5.138% | 10.28% | 15.42% | 456.6201 | 504.6853
Rb, ppm | 7.2499 | 0.64839 | 5.9531 | 8.5466 | 5.3047 | 9.195 | 8.944% | 17.89% | 26.83% | 6.8874 | 7.6124
S, wt.% | 2.948 | 0.17241 | 2.6031 | 3.2928 | 2.4307 | 3.4652 | 5.848% | 11.7% | 17.55% | 2.8006 | 3.0954
Sb, ppm | 324.4481 | 23.2822 | 277.8838 | 371.0124 | 254.6016 | 394.2945 | 7.176% | 14.35% | 21.53% | 308.2257 | 340.6705
Sc, ppm | 0.84326 | 0.079013 | 0.68524 | 1.0013 | 0.60622 | 1.0803 | 9.37% | 18.74% | 28.11% | 0.8011 | 0.88543
Se, ppm | 32.5281 | 4.0788 | 24.3706 | 40.6856 | 20.2918 | 44.7643 | 12.54% | 25.08% | 37.62% | 30.9017 | 34.1545
Sn, ppm | 29.5424 | 1.7744 | 25.9936 | 33.0912 | 24.2192 | 34.8656 | 6.006% | 12.01% | 18.02% | 28.0653 | 31.0195
Sr, ppm | 38.4178 | 5.5707 | 27.2763 | 49.5593 | 21.7056 | 55.13 | 14.5% | 29% | 43.5% | 36.4969 | 40.3387
Te, ppm | 48.9911 | 2.7419 | 43.5072 | 54.475 | 40.7653 | 57.217 | 5.597% | 11.19% | 16.79% | 46.5416 | 51.4407
Th, ppm | 3.1383 | 0.27289 | 2.5925 | 3.6841 | 2.3196 | 3.957 | 8.696% | 17.39% | 26.09% | 2.9814 | 3.2952
Tl, ppm | 1.8124 | 0.080446 | 1.6515 | 1.9733 | 1.5711 | 2.0538 | 4.439% | 8.877% | 13.32% | 1.7218 | 1.9031
U, ppm | 1.1214 | 0.091983 | 0.9374 | 1.3053 | 0.84541 | 1.3973 | 8.203% | 16.41% | 24.61% | 1.0653 | 1.1774
V, ppm | 12.6768 | 0.90776 | 10.8612 | 14.4923 | 9.9535 | 15.4 | 7.161% | 14.32% | 21.48% | 12.0429 | 13.3106
W, ppm | 4.3141 | 0.48061 | 3.3529 | 5.2753 | 2.8722 | 5.7559 | 11.14% | 22.28% | 33.42% | 4.0984 | 4.5298
Y, ppm | 3.0946 | 0.23655 | 2.6215 | 3.5677 | 2.3849 | 3.8043 | 7.644% | 15.29% | 22.93% | 2.9399 | 3.2493
Zn, ppm | 2057.7708 | 80.2389 | 1897.293 | 2218.2485 | 1817.0541 | 2298.4874 | 3.899% | 7.799% | 11.7% | 1954.8822 | 2160.6593
Zr, ppm | 12.5377 | 1.0477 | 10.4424 | 14.633 | 9.3947 | 15.6807 | 8.356% | 16.71% | 25.07% | 11.9108 | 13.1646

## Confidence & Tolerance Limits
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 611
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 15.6989 | 0.60102 | 15.4656 | 15.9322 | 15.6265 | 15.7714
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 15.5317 | 0.40669 | 15.2594 | 15.804 | 15.4532 | 15.6102
Infrared Combustion
S, Sulphur (wt.%) | 4.2615 | 0.17658 | 4.1812 | 4.3417 | 4.2059 | 4.317
4-Acid Digestion
Ag, Silver (ppm) | 80.0064 | 1.613 | 79.4686 | 80.5443 | 78.2261 | 81.7868
Al, Aluminium (wt.%) | 5.7836 | 0.21424 | 5.6877 | 5.8795 | 5.6575 | 5.9097
As, Arsenic (ppm) | 3399.6771 | 179.6305 | 3314.3745 | 3484.9797 | 3337.8255 | 3461.5288
Be, Beryllium (ppm) | 1.4613 | 0.1064 | 1.4094 | 1.5132 | 1.396 | 1.5266
Bi, Bismuth (ppm) | 265.4204 | 10.9916 | 260.6082 | 270.2326 | 259.7337 | 271.1071
Ca, Calcium (wt.%) | 0.22729 | 0.01767 | 0.21972 | 0.23485 | 0.21774 | 0.23684
Cd, Cadmium (ppm) | 13.912 | 0.73495 | 13.5795 | 14.2446 | 13.5302 | 14.2938
Ce, Cerium (ppm) | 46.4045 | 4.4604 | 43.9051 | 48.9039 | 44.3607 | 48.4483
Co, Cobalt (ppm) | 8.6962 | 0.26628 | 8.5891 | 8.8033 | 8.3859 | 9.0065
Cr, Chromium (ppm) | 50.5179 | 7.1553 | 47.4618 | 53.5741 | 47.9936 | 53.0423
Cs, Caesium (ppm) | 2.0619 | 0.14766 | 1.991 | 2.1327 | 1.9634 | 2.1603
Cu, Copper (wt.%) | 1.1667 | 0.021569 | 1.1583 | 1.1751 | 1.1494 | 1.1841
Dy, Dysprosium (ppm) | 1.5513 | 0.1096 | 1.4264 | 1.6762 | IND | IND
Er, Erbium (ppm) | 0.613 | 0.041514 | 0.58608 | 0.63992 | IND | IND
Eu, Europium (ppm) | 0.75567 | 0.054989 | 0.71096 | 0.80038 | IND | IND
Fe, Iron (wt.%) | 2.5366 | 0.092493 | 2.4981 | 2.5752 | 2.4872 | 2.5861
Ga, Gallium (ppm) | 24.439 | 1.6781 | 23.7038 | 25.1741 | 23.6255 | 25.2524
Gd, Gadolinium (ppm) | 2.9173 | 0.2339 | 2.6579 | 3.1767 | 2.5092 | 3.3254
Hf, Hafnium (ppm) | 2.2641 | 0.11733 | 2.2099 | 2.3184 | 2.1637 | 2.3645
Ho, Holmium (ppm) | 0.21867 | 0.023887 | 0.18679 | 0.25054 | IND | IND
In, Indium (ppm) | 4.6787 | 0.26784 | 4.5496 | 4.8079 | 4.5544 | 4.803
K, Potassium (wt.%) | 1.8594 | 0.069803 | 1.8304 | 1.8883 | 1.8237 | 1.8951
La, Lanthanum (ppm) | 20.7995 | 3.6259 | 19.0689 | 22.5302 | 20.125 | 21.474
Li, Lithium (ppm) | 29.6931 | 1.5518 | 29.0352 | 30.351 | 28.7781 | 30.608
Mg, Magnesium (ppm) | 1472.4424 | 93.1452 | 1433.5632 | 1511.3216 | 1430.19 | 1514.6948
Mn, Manganese (ppm) | 78.8582 | 3.7586 | 77.2758 | 80.4406 | 76.8803 | 80.8362
Mo, Molybdenum (ppm) | 5.4576 | 0.37019 | 5.303 | 5.6122 | 5.2307 | 5.6845
Na, Sodium (wt.%) | 0.80378 | 0.028317 | 0.79177 | 0.81578 | 0.78493 | 0.82263
Nb, Niobium (ppm) | 8.8228 | 0.56945 | 8.5411 | 9.1045 | 8.5322 | 9.1134
Nd, Neodymium (ppm) | 19.393 | 1.2368 | 18.037 | 20.749 | 18.051 | 20.735
Ni, Nickel (ppm) | 27.0308 | 1.0901 | 26.5796 | 27.4819 | 26.1452 | 27.9164
P, Phosphorus (ppm) | 547.8208 | 22.2632 | 537.5579 | 558.0837 | 535.1479 | 560.4937
Pb, Lead (ppm) | 639.5673 | 34.8959 | 624.4179 | 654.7167 | 625.4245 | 653.7101
Pr, Praseodymium (ppm) | 5.1463 | 0.58641 | 4.3979 | 5.8948 | 4.7788 | 5.5139
Rb, Rubidium (ppm) | 64.5088 | 2.9125 | 63.071 | 65.9466 | 62.9961 | 66.0215
S, Sulphur (wt.%) | 4.177 | 0.16284 | 4.1103 | 4.2437 | 4.1175 | 4.2365
Sb, Antimony (ppm) | 364.7518 | 30.63 | 351.3534 | 378.1502 | 357.1232 | 372.3804
Sc, Scandium (ppm) | 3.2 | 0.22418 | 3.0899 | 3.31 | 3.0408 | 3.3591
Se, Selenium (ppm) | 34.7806 | 3.4058 | 33.1486 | 36.4125 | 32.7572 | 36.8039
Sm, Samarium (ppm) | 3.8013 | 0.22495 | 3.5501 | 4.0526 | 3.6213 | 3.9813
Sn, Tin (ppm) | 32.0502 | 2.3449 | 31.0566 | 33.0439 | 31.2716 | 32.8289
Sr, Strontium (ppm) | 316.8633 | 21.3289 | 306.774 | 326.9527 | 308.9508 | 324.7759
Ta, Tantalum (ppm) | 0.71394 | 0.04496 | 0.68751 | 0.74038 | 0.6911 | 0.73679
Tb, Terbium (ppm) | 0.31079 | 0.065442 | 0.25673 | 0.36486 | IND | IND
Te, Tellurium (ppm) | 49.2816 | 2.7822 | 47.955 | 50.6082 | 47.961 | 50.6023
Th, Thorium (ppm) | 8.6627 | 1.0794 | 8.106 | 9.2194 | 8.2584 | 9.067
Ti, Titanium (wt.%) | 0.19187 | 0.006893 | 0.18888 | 0.19486 | 0.18852 | 0.19522
Tl, Thallium (ppm) | 2.1746 | 0.11418 | 2.1197 | 2.2295 | 2.1078 | 2.2414
U, Uranium (ppm) | 2.6771 | 0.16446 | 2.5985 | 2.7556 | 2.5907 | 2.7634
V, Vanadium (ppm) | 31.8691 | 1.498 | 31.2347 | 32.5035 | 30.9817 | 32.7565
W, Tungsten (ppm) | 8.7494 | 0.45404 | 8.5407 | 8.958 | 8.4245 | 9.0743
Y, Yttrium (ppm) | 6.8039 | 0.33754 | 6.6619 | 6.9459 | 6.5531 | 7.0547
Yb, Ytterbium (ppm) | 0.55042 | 0.077238 | 0.48956 | 0.61127 | IND | IND
Zn, Zinc (ppm) | 2022.9122 | 72.8504 | 1993.3525 | 2052.472 | 1984.8879 | 2060.9365
Zr, Zirconium (ppm) | 68.8064 | 5.3153 | 66.4525 | 71.1604 | 66.7414 | 70.8714
Aqua Regia Digestion
Ag, Silver (ppm) | 79.2496 | 3.6206 | 77.6291 | 80.87 | 77.847 | 80.6521
Al, Aluminium (wt.%) | 0.83827 | 0.04819 | 0.81523 | 0.86131 | 0.81654 | 0.86001
As, Arsenic (ppm) | 3340.4787 | 192.0963 | 3252.9229 | 3428.0345 | 3268.186 | 3412.7714
Be, Beryllium (ppm) | 0.28063 | 0.024551 | 0.26859 | 0.29266 | 0.26173 | 0.29952
Bi, Bismuth (ppm) | 256.488 | 15.897 | 249.2806 | 263.6954 | 250.2346 | 262.7414
Ca, Calcium (wt.%) | 0.10798 | 0.0048076 | 0.10644 | 0.10952 | 0.10588 | 0.11009
Cd, Cadmium (ppm) | 13.7293 | 0.97814 | 13.2833 | 14.1754 | 13.2727 | 14.186
Ce, Cerium (ppm) | 14.6378 | 1.45 | 13.7318 | 15.5437 | 14.2172 | 15.0583
Co, Cobalt (ppm) | 8.6376 | 0.46892 | 8.4269 | 8.8483 | 8.3062 | 8.9691
Cr, Chromium (ppm) | 41.6612 | 2.9096 | 40.3622 | 42.9601 | 40.5441 | 42.7782
Cs, Caesium (ppm) | 0.73252 | 0.07014 | 0.69064 | 0.77441 | 0.71287 | 0.75218
Cu, Copper (wt.%) | 1.1791 | 0.023749 | 1.1704 | 1.1877 | 1.1588 | 1.1993
Fe, Iron (wt.%) | 2.4047 | 0.1238 | 2.3508 | 2.4586 | 2.3616 | 2.4477
Ga, Gallium (ppm) | 6.83 | 0.50387 | 6.5577 | 7.1024 | 6.6286 | 7.0315
Hf, Hafnium (ppm) | 0.41875 | 0.030734 | 0.40198 | 0.43552 | 0.3978 | 0.4397
Hg, Mercury (ppm) | 0.94658 | 0.045683 | 0.92065 | 0.97252 | 0.9124 | 0.98076
In, Indium (ppm) | 4.5275 | 0.17201 | 4.4307 | 4.6243 | 4.3917 | 4.6632
K, Potassium (wt.%) | 0.2007 | 0.01152 | 0.19547 | 0.20593 | 0.193 | 0.2084
La, Lanthanum (ppm) | 6.8578 | 0.66678 | 6.508 | 7.2075 | 6.5981 | 7.1175
Li, Lithium (ppm) | 8.2534 | 0.61915 | 7.892 | 8.6147 | 7.9235 | 8.5832
Mg, Magnesium (ppm) | 999.8813 | 67.3214 | 973.1317 | 1026.631 | 943.6475 | 1056.1152
Mn, Manganese (ppm) | 61.9307 | 3.5685 | 60.3454 | 63.516 | 59.877 | 63.9844
Mo, Molybdenum (ppm) | 5.1262 | 0.3195 | 5.0013 | 5.251 | 4.9257 | 5.3266
Na, Sodium (wt.%) | 0.044336 | 0.0081792 | 0.040717 | 0.047956 | 0.041381 | 0.047291
Nb, Niobium (ppm) | 0.16821 | 0.030118 | 0.14535 | 0.19106 | IND | IND
Ni, Nickel (ppm) | 26.9125 | 1.7578 | 26.1425 | 27.6824 | 25.7251 | 28.0998
P, Phosphorus (ppm) | 242.618 | 12.1583 | 236.1085 | 249.1275 | 233.1386 | 252.0973
Pb, Lead (ppm) | 480.6527 | 24.6983 | 469.952 | 491.3534 | 471.2332 | 490.0722
Rb, Rubidium (ppm) | 7.2499 | 0.64839 | 6.8431 | 7.6566 | 6.9847 | 7.5151
S, Sulphur (wt.%) | 2.948 | 0.17241 | 2.8673 | 3.0286 | 2.8906 | 3.0053
Sb, Antimony (ppm) | 324.4481 | 23.2822 | 313.2862 | 335.61 | 316.9051 | 331.9911
Sc, Scandium (ppm) | 0.84326 | 0.079013 | 0.79929 | 0.88724 | IND | IND
Se, Selenium (ppm) | 32.5281 | 4.0788 | 30.4243 | 34.6319 | 31.0975 | 33.9586
Sn, Tin (ppm) | 29.5424 | 1.7744 | 28.7565 | 30.3284 | 28.9538 | 30.131
Sr, Strontium (ppm) | 38.4178 | 5.5707 | 35.6928 | 41.1428 | 37.5139 | 39.3216
Te, Tellurium (ppm) | 48.9911 | 2.7419 | 47.6372 | 50.3451 | 47.5374 | 50.4449
Th, Thorium (ppm) | 3.1383 | 0.27289 | 2.9869 | 3.2897 | 3.0186 | 3.258
Tl, Thallium (ppm) | 1.8124 | 0.080446 | 1.7731 | 1.8517 | 1.7426 | 1.8823
U, Uranium (ppm) | 1.1214 | 0.091983 | 1.0744 | 1.1683 | 1.0873 | 1.1554
V, Vanadium (ppm) | 12.6768 | 0.90776 | 12.2742 | 13.0793 | 12.2206 | 13.1329
W, Tungsten (ppm) | 4.3141 | 0.48061 | 4.0572 | 4.5709 | 4.166 | 4.4621
Y, Yttrium (ppm) | 3.0946 | 0.23655 | 2.981 | 3.2082 | 2.9884 | 3.2008
Zn, Zinc (ppm) | 2057.7708 | 80.2389 | 2025.2139 | 2090.3277 | 2016.0374 | 2099.5041
Zr, Zirconium (ppm) | 12.5377 | 1.0477 | 12.0213 | 13.0541 | 12.2124 | 12.863

## Indicative Values
Table 3. Indicative Values for OREAS 611
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 5 | Pt | ppb | < 5
Infrared Combustion
C | wt.% | 0.07
4-Acid Digestion
Ba | ppm | 263.0879 | Hg | ppm | 0.655 | Re | ppb | 2.0741
Ge | ppm | 2.3539 | Lu | ppb | 86.6667 | Tm | ppb | 91.6667
Aqua Regia Digestion
B | ppm | < 10 | Ho | ppm | 0.1 | Sm | ppm | 1.4
Ba | ppm | 57.7586 | Lu | ppb | 31.6667 | Ta | ppm | < 0.01
Dy | ppm | 0.83333 | Nd | ppm | 8.325 | Tb | ppm | 0.16917
Er | ppm | 0.28333 | Pd | ppb | 195.8333 | Ti | wt.% | 0.010127
Eu | ppm | 0.21667 | Pr | ppm | 2.05 | Tm | ppb | < 100
Gd | ppm | 1.4667 | Pt | ppb | < 1 | Yb | ppm | 0.2
Ge | ppm | 0.32071 | Re | ppb | 1.8333
Borate Fusion XRF
Al2O3 | wt.% | 11.28 | Fe2O3 | wt.% | 3.62 | S | wt.% | 4.3268
As | ppm | 3635 | K2O | wt.% | 2.29 | SiO2 | wt.% | 69.535
BaO | ppm | 5075 | MgO | wt.% | 0.275 | Sn | ppm | 20
CaO | wt.% | 0.31 | MnO | wt.% | 0.012 | Sr | ppm | 338.2378
Cl | ppm | 30 | Na2O | wt.% | 1.13 | TiO2 | wt.% | 0.323
Co | ppm | 15 | Ni | ppm | 35 | V2O5 | ppm | 55
Cr2O3 | ppm | 90 | P2O5 | wt.% | 0.1285 | Zn | ppm | 2080
Cu | wt.% | 1.165 | Pb | ppm | 690 | Zr | ppm | 196.1853
Thermogravimetry
LOI1000 | wt.% | 8.48
Peroxide Fusion ICP
Cu | wt.% | 1.1483
Laser Ablation ICP-MS
Ag | ppm | 96.25 | Hf | ppm | 5.215 | Sm | ppm | 4.09
As | ppm | 3550 | Ho | ppm | 0.255 | Sn | ppm | 32.1
Ba | ppm | 4445 | In | ppm | 4.3 | Sr | ppm | 349
Be | ppm | 1.7 | La | ppm | 26.55 | Ta | ppm | 0.77
Bi | ppm | 259.5 | Lu | ppb | 105 | Tb | ppm | 0.38
Cd | ppm | 16.75 | Mn | ppm | 75 | Te | ppm | 53
Ce | ppm | 50.5 | Mo | ppm | 5.2 | Th | ppm | 10.35
Co | ppm | 8.55 | Nb | ppm | 9.235 | Ti | wt.% | 0.192
Cr | ppm | 60.5 | Nd | ppm | 21.1 | Tl | ppm | 2.9
Cs | ppm | 2.005 | Ni | ppm | 32 | Tm | ppb | 100
Cu | wt.% | 1.165 | Pb | ppm | 698.5 | U | ppm | 2.755
Dy | ppm | 1.66 | Pr | ppm | 5.69 | V | ppm | 33
Er | ppm | 0.715 | Rb | ppm | 61.1 | W | ppm | 8.75
Eu | ppm | 0.79 | Re | ppb | 17.5 | Y | ppm | 7.6
Ga | ppm | 23.3 | Sb | ppm | 386.5 | Yb | ppm | 0.72
Gd | ppm | 3.015 | Sc | ppm | 3.75 | Zn | ppm | 2015
Ge | ppm | 5.125 | Se | ppm | < 5 | Zr | ppm | 178.5
X-ray Photon Assay
Au | ppm | 15.808

## Fire Assay
| Analytical results for Au in OREAS 611 (Certified Value 15.7 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*GRAV | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*GRAV | FA*GRAV | FA*OES | FA*OES | FA*AAS | FA*GRAV | FA*AAS | FA*GRAV | FA*AAS | FA*AAS | FA*GRAV | FA*AAS | FA*GRAV | 2
|  |  | 0.085g | 25g | 25g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 40g | 30g | 3
| 1 | 1 | 16.0738 | 14.9 | 16 | 15.7 | 14.5 | 15.5 | 17.25 | 16.25 | 15.85 | 15.95 | 16.7 | 15.437 | 15.91 | 16.2 | 16.1 | 15.14 | 15.9 | 15.774 | 15.76 | 15.4 | 15.82 | 16.2 | 15.46 | 15.2 | 15.52 | 15.57 | 1
| 1 | 2 | 16.8154 | 14.9 | 15.987 | 16.6 | 14.3 | 15.3 | 16.8 | 15.85 | 15.45 | 15.6 | 16.6 | 15.478 | 15.93 | 16.4 | 16.6 | 15.32 | 15.4 | 15.27 | 15.6 | 15.5 | 15.78 | 16.35 | 15.7 | 14.5 | 15.56 | 15.16 | 47
| 1 | 3 | 16.2295 | 15.1 | 16.099 | 15.7 | 14.7 | 15.35 | 16.85 | 16.3 | 15.75 | 15.45 | 16.5 | 15.658 | 15.96 | 16.3 | 16.5 | 15.7 | 15.9 | 15.2273 | 15.255 | 15.6 | 15.98 | 15.93 | 15.98 | 14.7 | 15.48 | 15 | 16
| 1 | 4 | 16.2144 | 15 | 16.242 | 15.7 | 14.4 | 15.65 | 16.4 | 15.85 | 16.1 | 15.45 | 16.7 | 15.987 | 15.95 | 16.1 | 16.9 | 15.1 | 15.2 | 15.162 | 15.535 | 15.1 | 15.29 | 16.86 | 16.6 | 15 | 15.52 | 15.22 | 15.6989
| 1 | 5 | 16.1665 | 15.3 | 15.832 | 15.9 | 14.7 | 15.25 | 16.6 | 16.15 | 15.8 | 15.2 | 16.4 | 15.652 | 15.9433 | 16.2 | 16.2 | 15.1 | 15.9 | 15.3013 | 15.29 | 15.1 | 15.68 | 17.32 | 15.8 | 15.1 | 15.59 | 14.995 | 7
| 1 | 6 | 16.1405 | 15.4 | 15.59 | 16 | 14.4 | 15.4 | 16.9 | 16.5 | 15.15 | 15.5 | 16.5 | 15.852 | 15.92 | 16.2 | 16.7 | 15.2 | 15.8 | 15.434 | 15.485 | 15.2 | 15.79 | 17.04 | 15.66 | 14.7 | 15.57 | 14.914
|  | 7 | 00: 15.9103
|  | 8 | 00: 17.5504
|  | 9 | 00: 16.4625
|  | 10 | 00: 15.8244
|  | 11 | 00: 16.1553
|  | 12 | 00: 15.7382
|  | 13 | 00: 16.6956
|  | 14 | 00: 16.2763
|  | 15 | 00: 15.785
|  | 16 | 00: 16.7516
|  | 17 | 00: 16.7733
|  | 18 | 00: 15.8072
|  | 19 | 00: 16.3264
|  | 20 | 00: 16.1926
| Mean |  | 16.2945 | 15.1 | 15.9583 | 15.9333 | 14.5 | 15.4083 | 16.8 | 16.15 | 15.6833 | 15.525 | 16.5667 | 15.6773 | 15.9356 | 16.2333 | 16.5 | 15.26 | 15.6833 | 15.3614 | 15.4875 | 15.3167 | 15.7233 | 16.6167 | 15.8667 | 14.8667 | 15.54 | 15.1432
| Median |  | 16.2035 | 15.05 | 15.9935 | 15.8 | 14.45 | 15.375 | 16.825 | 16.2 | 15.775 | 15.475 | 16.55 | 15.655 | 15.9367 | 16.2 | 16.55 | 15.17 | 15.85 | 15.2857 | 15.51 | 15.3 | 15.785 | 16.605 | 15.75 | 14.85 | 15.54 | 15.08
| Std Dev. |  | 0.44612 | 0.20976 | 0.22558 | 0.35024 | 0.16733 | 0.14634 | 0.2881 | 0.25884 | 0.33417 | 0.24648 | 0.12111 | 0.21209 | 0.018935 | 0.10328 | 0.30332 | 0.23082 | 0.30605 | 0.22145 | 0.19091 | 0.2137 | 0.23347 | 0.53839 | 0.39773 | 0.27325 | 0.040497 | 0.23805
| Rel.Std.Dev. |  | 2.738% | 1.389% | 1.414% | 2.198% | 1.154% | 0.9498% | 1.715% | 1.603% | 2.131% | 1.588% | 0.731% | 1.353% | 0.1188% | 0.6362% | 1.838% | 1.513% | 1.951% | 1.442% | 1.233% | 1.395% | 1.485% | 3.24% | 2.507% | 1.838% | 0.2606% | 1.572%
| PDM3 |  | 3.794% | -3.815% | 1.652% | 1.493% | -7.637% | -1.851% | 7.014% | 2.873% | -0.09924% | -1.108% | 5.527% | -0.1375% | 1.507% | 3.404% | 5.103% | -2.796% | -0.09924% | -2.15% | -1.347% | -2.435% | 0.1556% | 5.846% | 1.069% | -5.301% | -1.012% | -3.54%
| Z-Score (Absolute) |  | NA | 1.22 | 0.58 | 0.52 | 2.48 | 0.58 | 2.34 | 0.98 | 0 | 0.33 | 1.85 | 0.01 | 0.53 | 1.15 | 1.71 | 0.89 | 0 | 0.67 | 0.41 | 0.77 | 0.08 | 1.96 | 0.38 | 1.71 | 0.3 | 1.13
| Analytical results for Pd in OREAS 611 (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 00: Lab | 1
| No | No | 00: 2.16 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 40g | 2
| 1 | 1 | 00: < 5 | 1
| 1 | 2 | 00: < 5 | 1
| 1 | 3 | 00: < 5 | 16
| 1 | 4 | 00: < 5 | < 5
| 1 | 5 | 00: < 5 | 7
| 1 | 6 | 00: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00
| Analytical results for Pt in OREAS 611 (Indicative Value < 5 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 00: Lab | 1
| No | No | 00: 2.16 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 40g | 2
| 1 | 1 | 00: < 5 | 1
| 1 | 2 | 00: < 5 | 1
| 1 | 3 | 00: < 5 | 16
| 1 | 4 | 00: < 5 | < 5
| 1 | 5 | 00: < 5 | 7
| 1 | 6 | 00: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 611 (Certified Value 15.53 ppm) | [16 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.16 | 2.24 | 2.25 | ppm
|  |  | INAA | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | 2
|  |  | 0.085g | 15g | 30g | 25g | 25g | 25g | 25g | 25g | 30g | 25g | 25g | 15g | 40g | 30g | 25g | 3
| 1 | 1 | 16.0738 | 17.2839 | 16 | 16.15 | 16.2 | 16.45 | 15.25 | 14.95 | 15.6 | > 15 | 15.2476 | > 10 | 15 | 14.7 | 15.65 | 1
| 1 | 2 | 16.8154 | 16.9381 | 16 | 15.55 | 16.25 | 17 | 15.95 | 15.5 | 15.5 | > 15 | 15.5333 | > 10 | 15.3 | 14.61 | 15.5 | 47
| 1 | 3 | 16.2295 | 17.2356 | 16 | 15.8 | 16.25 | 16.15 | 15.85 | 14.9 | 15.6 | > 15 | 15.3737 | > 10 | 15.8 | 14.58 | 15.35 | 16
| 1 | 4 | 16.2144 | 17.1318 | 14.9 | 15.95 | 16.4 | 16.3 | 15.35 | 14.9 | 15.4 | > 15 | 15.3599 | > 10 | 15.5 | 14.61 | 15.2 | 15.5317
| 1 | 5 | 16.1665 | 17.8768 | 15.6 | 15.65 | 16 | 16.35 | 15.45 | 15.1 | 15.3 | > 15 | 15.4916 | > 10 | 15.5 | 14.58 | 15.16 | 9
| 1 | 6 | 16.1405 | 17.0798 | 15.3 | 15.6 | 16.1 | 16.6 | 16.1 | 14.7 | 15.2 | > 15 | 15.2361 | > 10 | 15.7 | 14.61 | 15.1
|  | 7 | 00: 15.9103
|  | 8 | 00: 17.5504
|  | 9 | 00: 16.4625
|  | 10 | 00: 15.8244
|  | 11 | 00: 16.1553
|  | 12 | 00: 15.7382
|  | 13 | 00: 16.6956
|  | 14 | 00: 16.2763
|  | 15 | 00: 15.785
|  | 16 | 00: 16.7516
|  | 17 | 00: 16.7733
|  | 18 | 00: 15.8072
|  | 19 | 00: 16.3264
|  | 20 | 00: 16.1926
| Mean |  | 16.2945 | 17.2577 | 15.6333 | 15.7833 | 16.2 | 16.475 | 15.6583 | 15.0083 | 15.4333 |  | 15.3737 |  | 15.4667 | 14.615 | 15.3267
| Median |  | 16.2035 | 17.1837 | 15.8 | 15.725 | 16.225 | 16.4 | 15.65 | 14.925 | 15.45 |  | 15.3668 |  | 15.5 | 14.61 | 15.275
| Std Dev. |  | 0.44612 | 0.32683 | 0.45898 | 0.23166 | 0.13784 | 0.29791 | 0.35273 | 0.27279 | 0.1633 |  | 0.12197 |  | 0.28752 | 0.044159 | 0.21463
| Rel.Std.Dev. |  | 2.738% | 1.894% | 2.936% | 1.468% | 0.8509% | 1.808% | 2.253% | 1.818% | 1.058% |  | 0.7934% |  | 1.859% | 0.3021% | 1.4%
| PDM3 |  | 4.911% | 11.11% | 0.6543% | 1.62% | 4.303% | 6.073% | 0.8153% | -3.37% | -0.6334% |  | -1.017% |  | -0.4188% | -5.902% | -1.32%
| Z-Score (Absolute) |  | NA | 5.04 | 0.25 | 0.69 | 1.92 | 2.73 | 0.32 | 1.6 | 0.34 | NA | 0.52 | NA | 0.25 | 2.76 | 0.66
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 611 (Indicative Value 0.07 wt.%) | [22 empty] | Indicative
C | Round | Replicate | BV Geo | [20 empty] | 1
| No | No | Perth | [20 empty] | wt.%
|  |  | IRC | [20 empty] | 3
3
| 1 | 1 | 0.07 | [20 empty] | 1
| 1 | 2 | 0.07 | [20 empty] | 3
| Mean |  | 0.07 | [20 empty] | 16
| Median |  | 0.07 | [20 empty] | 0.07
| Std Dev. |  | 0 | [20 empty] | 9
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 611 (Certified Value 4.26 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 4.17 | 4.05 | 4.31 | 4.45 | 4.3 | 4.24 | 4.22 | 4.33 | 4.14 | 4.39 | 4.4537 | 4.3 | 4.5 | 4.06 | 4.48 | 4.01 | 4.32 | 4.31 | 3.96 | 4.81 | 3.965 | 1
| 1 | 2 | 4.15 | 3.98 | 4.28 | 4.52 | 4.38 | 4.18 | 4.22 | 4.63 | 4.14 | 4.34 | 4.4045 | 4.29 | 4.44 | 4.07 | 4.49 | 4.03 | 4.31 | 4.4 | 4.05 | 4.85 | 3.99 | 16
| 1 | 3 | 4.17 | 3.97 | 4.18 | 4.52 | 4.46 | 4.13 | 4.26 | 4.7 | 4.14 | 4.34 | 4.42 | 4.31 | 4.39 | 4.07 | 4.5 | 4.04 | 4.27 | 4.41 | 4.04 | 4.84 | 3.994 | 16
| 1 | 4 | 4.15 | 4.02 | 4.18 | 4.5 | 4.43 | 4.19 | 4.24 | 4.36 | 4.14 | 4.36 | 4.4306 | 4.27 | 4.48 | 4.07 | 4.45 | 3.98 | 4.33 | 4.36 | 4.11 | 4.86 | 4.0585 | 4.2615
| 1 | 5 | 4.21 | 4.14 | 4.34 | 4.43 | 4.48 | 4.2 | 4.17 | 4.69 | 4.15 | 4.26 | 4.4307 | 4.28 | 4.39 | 4 | 4.46 | 3.98 | 4.3 | 4.38 | 3.49 | 4.86 | 4.1203 | 11
| 1 | 6 | 4.17 | 3.88 | 4.41 | 4.38 | 4.47 | 4.13 | 4.21 | 4.49 | 4.14 | 4.35 | 4.4525 | 4.32 | 4.43 | 4.07 | 4.49 | 4 | 4.25 | 4.31 | 4.18 | 4.79 | 4.0875
| Mean |  | 4.17 | 4.0067 | 4.2833 | 4.4667 | 4.42 | 4.1783 | 4.22 | 4.5333 | 4.1417 | 4.34 | 4.432 | 4.295 | 4.4383 | 4.0567 | 4.4783 | 4.0067 | 4.2967 | 4.3617 | 3.9717 | 4.835 | 4.0359
| Median |  | 4.17 | 4 | 4.295 | 4.475 | 4.445 | 4.185 | 4.22 | 4.56 | 4.14 | 4.345 | 4.4306 | 4.295 | 4.435 | 4.07 | 4.485 | 4.005 | 4.305 | 4.37 | 4.045 | 4.845 | 4.0263
| Std Dev. |  | 0.021909 | 0.087101 | 0.090921 | 0.056451 | 0.068993 | 0.042622 | 0.030332 | 0.16428 | 0.0040825 | 0.043359 | 0.01894 | 0.018708 | 0.04535 | 0.028048 | 0.019408 | 0.025033 | 0.030768 | 0.043551 | 0.24718 | 0.02881 | 0.061945
| Rel.Std.Dev. |  | 0.5254% | 2.174% | 2.123% | 1.264% | 1.561% | 1.02% | 0.7188% | 3.624% | 0.09857% | 0.9991% | 0.4274% | 0.4356% | 1.022% | 0.6914% | 0.4334% | 0.6248% | 0.7161% | 0.9985% | 6.224% | 0.5959% | 1.535%
| PDM3 |  | -2.146% | -5.979% | 0.5133% | 4.815% | 3.72% | -1.951% | -0.9729% | 6.38% | -2.811% | 1.843% | 4.002% | 0.7871% | 4.151% | -4.806% | 5.089% | -5.979% | 0.8262% | 2.351% | -6.8% | 13.46% | -5.293%
| Z-Score (Absolute) |  | 0.59 | 1.36 | 0.05 | 0.81 | 0.59 | 0.55 | 0.35 | 1.12 | 0.72 | 0.21 | 0.64 | 0 | 0.67 | 1.12 | 0.86 | 1.36 | 0.01 | 0.31 | 1.52 | 2.54 | 1.22
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 611 (Certified Value 80 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*AAS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 81.1 | 79.064 | 73.2 | 76.9 | 78.6 | 80.6 | 77.7 | 79 | 81.3 | 79.53 | 82.2065 | 85.06 | 74.2 | 80.5 | 81.3202 | 78.8 | 82 | > 10 | 77.65 | 77 | 70.5 | 81 | 78.44 | 1
| 1 | 2 | 79.3 | 78.219 | 74.2 | 77.5 | 78.1 | 78.8 | 76.7 | 80.1 | 80.8 | 81.58 | 80.0579 | 74.62 | 74.4 | 80 | 81.0465 | 79.4 | 80 | > 10 | 80.37 | 75 | 72.2 | 82 | 80.29 | 29
| 1 | 3 | 81.2 | 79.087 | 73.1 | 79.1 | 78.3 | 78.7 | 78.5 | 79.3 | 80.7 | 81.74 | 82.4742 | 77.37 | 73.2 | 80.5 | 81.9274 | 79.7 | 81 | > 10 | 79.88 | 77 | 74 | 82 | 79.14 | 16
| 1 | 4 | 81 | 79.688 | 74.3 | 78.6 | 79 | 79.3 | 79.5 | 79.8 | 80.3 | 79.84 | 81.339 | 83.28 | 74.4 | 79 | 81.1895 | 79.7 | 80 | > 10 | 81.47 | 76 | 68.8 | 82 | 80.99 | 80.0064
| 1 | 5 | 79 | 80.727 | 77.7 | 80.6 | 77.9 | 78.4 | 77.3 | 82.7 | 81 | 81.01 | 81.5679 | 71.23 | 74.4 | 80.5 | 80.2339 | 79.8 | 81 | > 10 | 78.39 | 75 | 68.1 | 83 | 81.84 | 13
| 1 | 6 | 81.3 | 75.217 | 77.2 | 78.1 | 78.3 | 80.1 | 76.7 | 81.3 | 80.2 | 81.06 | 80.4211 | 87.34 | 73.9 | 79 | 80.4053 | 80.5 | 79 | > 10 | 80.86 | 76 | 68 | 82 | 80.69
| Mean |  | 80.4833 | 78.667 | 74.95 | 78.4667 | 78.3667 | 79.3167 | 77.7333 | 80.3667 | 80.7167 | 80.7933 | 81.3444 | 79.8167 | 74.0833 | 79.9167 | 81.0205 | 79.65 | 80.5 |  | 79.77 | 76 | 70.2667 | 82 | 80.2317
| Median |  | 81.05 | 79.0755 | 74.25 | 78.35 | 78.3 | 79.05 | 77.5 | 79.95 | 80.75 | 81.035 | 81.4534 | 80.325 | 74.3 | 80.25 | 81.118 | 79.7 | 80.5 |  | 80.125 | 76 | 69.65 | 82 | 80.49
| Std Dev. |  | 1.042 | 1.8827 | 2.0047 | 1.3033 | 0.38816 | 0.86583 | 1.0985 | 1.3938 | 0.41673 | 0.90972 | 0.95678 | 6.3688 | 0.47504 | 0.73598 | 0.6227 | 0.55408 | 1.0488 |  | 1.473 | 0.89443 | 2.4378 | 0.63246 | 1.2472
| Rel.Std.Dev. |  | 1.295% | 2.393% | 2.675% | 1.661% | 0.4953% | 1.092% | 1.413% | 1.734% | 0.5163% | 1.126% | 1.176% | 7.979% | 0.6412% | 0.9209% | 0.7686% | 0.6956% | 1.303% |  | 1.847% | 1.177% | 3.469% | 0.7713% | 1.554%
| PDM3 |  | 0.5961% | -1.674% | -6.32% | -1.925% | -2.05% | -0.8622% | -2.841% | 0.4502% | 0.8877% | 0.9835% | 1.672% | -0.2372% | -7.403% | -0.1122% | 1.267% | -0.4455% | 0.6169% |  | -0.2955% | -5.008% | -12.17% | 2.492% | 0.2815%
| Z-Score (Absolute) |  | 0.44 | 0.71 | 3.07 | 0.84 | 0.9 | 0.3 | 1.31 | 0.36 | 0.59 | 0.63 | 0.98 | 0.01 | 3.62 | 0.08 | 0.78 | 0.09 | 0.45 | NA | 0.01 | 2.41 | 6.04 | 1.4 | 0.28
| Analytical results for Al in OREAS 611 (Certified Value 5.78 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 5.8629 | 5.62 | 5.73 | 5.69 | 5.58 | 5.78 | 5.5 | 5.71 | 5.9 | 5.8554 | 5.8213 | 4.7607 | 6.05 | 5.97 | 6.0473 | 5.98 | 4.592 | 4.84 | 5.78 | 5.63 | 6.09 | 2.75 | 5.692 | 1
| 1 | 2 | 5.7197 | 5.59 | 5.93 | 5.73 | 5.39 | 5.69 | 5.37 | 5.79 | 5.82 | 5.9869 | 5.8646 | 5.1799 | 6.1 | 5.95 | 5.9846 | 6.18 | 4.322 | 4.71 | 5.83 | 5.47 | 5.86 | 2.77 | 5.638 | #N/A
| 1 | 3 | 5.8076 | 5.57 | 5.89 | 5.85 | 5.33 | 5.65 | 5.63 | 5.72 | 5.83 | 6.0634 | 6.0224 | 5.3394 | 6.02 | 6 | 5.9733 | 6.25 | 4.642 | 4.69 | 5.85 | 5.61 | 5.86 | 2.8 | 5.664 | 16
| 1 | 4 | 5.7119 | 5.6 | 5.99 | 5.75 | 5.58 | 5.67 | 5.69 | 5.76 | 5.99 | 5.7862 | 5.7446 | 4.9719 | 6.11 | 6 | 5.9416 | 6.01 | 4.659 | 4.75 | 5.94 | 5.55 | 5.57 | 2.8 | 5.69 | 5.7836
| 1 | 5 | 5.8134 | 5.73 | 6.2 | 5.97 | 5.53 | 5.56 | 5.6 | 6 | 6.01 | 5.9564 | 5.9649 | 4.9961 | 6.01 | 5.93 | 6.0132 | 5.95 | 4.385 | 4.86 | 5.89 | 5.34 | 5.55 | 2.8 | 5.616 | 14
| 1 | 6 | 5.8597 | 5.33 | 6.32 | 5.78 | 5.43 | 5.69 | 5.5 | 5.85 | 5.74 | 5.8542 | 5.8757 | 5.4699 | 6.06 | 6.02 | 6.0851 | 5.92 | 4.337 | 4.62 | 5.92 | 5.68 | 5.71 | 2.79 | 5.642
| Mean |  | 5.7959 | 5.5733 | 6.01 | 5.795 | 5.4733 | 5.6733 | 5.5483 | 5.805 | 5.8817 | 5.9171 | 5.8822 | 5.1197 | 6.0583 | 5.9783 | 6.0075 | 6.0483 | 4.4895 | 4.745 | 5.8683 | 5.5467 | 5.7733 | 2.785 | 5.657
| Median |  | 5.8105 | 5.595 | 5.96 | 5.765 | 5.48 | 5.68 | 5.55 | 5.775 | 5.865 | 5.9059 | 5.8701 | 5.088 | 6.055 | 5.985 | 5.9989 | 5.995 | 4.4885 | 4.73 | 5.87 | 5.58 | 5.785 | 2.795 | 5.653
| Std Dev. |  | 0.066129 | 0.13186 | 0.21531 | 0.10114 | 0.1052 | 0.071181 | 0.11479 | 0.10821 | 0.10496 | 0.10268 | 0.099463 | 0.26097 | 0.040702 | 0.034303 | 0.052294 | 0.13438 | 0.15794 | 0.091815 | 0.059805 | 0.12437 | 0.2052 | 0.020736 | 0.03043
| Rel.Std.Dev. |  | 1.141% | 2.366% | 3.583% | 1.745% | 1.922% | 1.255% | 2.069% | 1.864% | 1.785% | 1.735% | 1.691% | 5.097% | 0.6718% | 0.5738% | 0.8705% | 2.222% | 3.518% | 1.935% | 1.019% | 2.242% | 3.554% | 0.7446% | 0.5379%
| PDM3 |  | 0.2126% | -3.635% | 3.915% | 0.1976% | -5.364% | -1.906% | -4.067% | 0.3705% | 1.696% | 2.308% | 1.706% | -11.48% | 4.751% | 3.368% | 3.872% | 4.578% | -22.37% | -17.96% | 1.466% | -4.096% | -0.177% | -51.85% | -2.188%
| Z-Score (Absolute) |  | 0 | 0.7 | 0.68 | 0 | 1.02 | 0.39 | 0.78 | 0.03 | 0.28 | 0.39 | 0.28 | 2.14 | 0.84 | 0.58 | 0.67 | 0.8 | 4.14 | 3.33 | 0.23 | 0.79 | 0.07 | 9.55 | 0.44
| Analytical results for As in OREAS 611 (Certified Value 3400 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 3341 | 2625.6 | 2970 | 3510 | 3430 | 3580 | 3370 | 3580 | 3190 | 3351.7 | 3427.6424 | 3394.5 | 3390 | 3820 | 3457.1521 | 3435 | 3452.678 | 3225 | 3080 | 3333 | 943.6 | 2954 | 2761 | 1
| 1 | 2 | 3287 | 2569.6 | 3060 | 3560 | 3390 | 3520 | 3460 | 3630 | 3140 | 3411 | 3513.4943 | 3352.9 | 3400 | 3820 | 3460.5425 | 3437 | 3377.055 | 3202 | 3049 | 3243 | > 1000 | 2900 | 2043 | 17
| 1 | 3 | 3403 | 2632.3 | 3080 | 3630 | 3450 | 3510 | 3460 | 3600 | 3120 | 3425.4 | 3599.4118 | 3354.7 | 3400 | 3770 | 3439.3979 | 3478 | 3464.925 | 3186 | 3066 | 3321 | > 1000 | 2908 | 2758 | 16
| 1 | 4 | 3294 | 2552.1 | 3130 | 3570 | 3430 | 3520 | 3480 | 3620 | 3260 | 3351.1 | 3475.8605 | 3339 | 3410 | 3770 | 3441.5037 | 3429 | 3430.759 | 3114 | 3055 | 3420 | > 1000 | 2926 | 2801 | 3399.6771
| 1 | 5 | 3357 | 2625.6 | 3250 | 3710 | 3440 | 3420 | 3430 | 3770 | 3170 | 3412.4 | 3564.1405 | 3354.9 | 3350 | 3730 | 3442.5029 | 3465 | 3471.264 | 3143 | 3067 | 3310 | > 1000 | 2860 | 2709 | 15
| 1 | 6 | 3397 | 2349.5 | 3220 | 3560 | 3500 | 3490 | 3430 | 3700 | 3250 | 3397.2 | 3449.9876 | 3503.4 | 3490 | 3730 | 3494.0132 | 3487 | 3431.861 | 3145 | 3065 | 3305 | > 1000 | 2920 | 2795
| Mean |  | 3346.5 | 2559.1167 | 3118.3333 | 3590 | 3440 | 3506.6667 | 3438.3333 | 3650 | 3188.3333 | 3391.4667 | 3505.0895 | 3383.2333 | 3406.6667 | 3773.3333 | 3455.852 | 3455.1667 | 3438.0903 | 3169.1667 | 3063.6667 | 3322 | 943.6 | 2911.3333 | 2644.5
| Median |  | 3349 | 2597.6 | 3105 | 3565 | 3435 | 3515 | 3445 | 3625 | 3180 | 3404.1 | 3494.6774 | 3354.8 | 3400 | 3770 | 3449.8275 | 3451 | 3442.2695 | 3165.5 | 3065.5 | 3315.5 | 943.6 | 2914 | 2759.5
| Std Dev. |  | 49.3548 | 107.9761 | 104.5785 | 70.1427 | 35.7771 | 52.0256 | 38.6868 | 71.5542 | 57.0672 | 32.2954 | 66.8475 | 61.7424 | 45.8984 | 40.332 | 20.6629 | 24.7096 | 34.2089 | 41.9305 | 10.7269 | 57.2713 |  | 31.2581 | 296.4994
| Rel.Std.Dev. |  | 1.475% | 4.219% | 3.354% | 1.954% | 1.04% | 1.484% | 1.125% | 1.96% | 1.79% | 0.9523% | 1.907% | 1.825% | 1.347% | 1.069% | 0.5979% | 0.7152% | 0.995% | 1.323% | 0.3501% | 1.724% |  | 1.074% | 11.21%
| PDM3 |  | -1.564% | -24.72% | -8.276% | 5.598% | 1.186% | 3.147% | 1.137% | 7.363% | -6.217% | -0.2415% | 3.101% | -0.4837% | 0.2056% | 10.99% | 1.652% | 1.632% | 1.13% | -6.78% | -9.884% | -2.285% | -72.24% | -14.36% | -22.21%
| Z-Score (Absolute) |  | 0.26 | 4.87 | 1.6 | 1.16 | 0.28 | 0.67 | 0.27 | 1.51 | 1.19 | 0 | 0.67 | 0.05 | 0.09 | 2.24 | 0.38 | 0.37 | 0.27 | 1.3 | 1.92 | 0.41 | 14.33 | 2.81 | 4.37
| Analytical results for Ba in OREAS 611 (Indicative Value 263 ppm) | [24 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.12 | 2.14 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 764 | 27 | 42 | 160 | 110 | 1140 | 440 | 280 | 359 | 171.4848 | 94 | 3390 | 367.7696 | 2783 | 200.384 | 439 | 150 | 109 | > 1000 | 232 | 3716 |  |  | 1
| 1 | 2 | 845 | 37 | 43 | 170 | 100 | 990 | 290 | 320 | 357 | 176.3202 | 256 | 3420 | 410.8186 | 2556 | 203.52 | 433 | 54 | 116 | > 1000 | 235 | 3771 |  |  | 5
| 1 | 3 | 1127 | 46 | 36 | 190 | 70 | 740 | 840 | 320 | 384 | 169.5655 | 264 | 2770 | 398.8764 | 2621 | 199.599 | 441 | 75 | 92 | > 1000 | 235 | 3744 |  |  | 16
| 1 | 4 | 1324 | 29 | 29 | 170 | 80 | 1140 | 960 | 260 | 383 | 170.4242 | 245 | 2360 | 326.0922 | 2753 | 201.104 | 434 | 62 | 88 | > 1000 | 237 | 3765 |  |  | 263.0879
| 1 | 5 | 1662 | 24 | 43 | 200 | 120 | 630 | 830 | 400 | 382 | 162.5905 | 239 | 2350 | 315.6937 | 2703 | 198.499 | 447 | 75 | 96 | > 1000 | 234 | 3801 |  |  | 11
| 1 | 6 | 1385 | 45 | 42 | 230 | 360 | 530 | 710 | 260 | 397 | 166.0913 | 90 | 2330 | 336.1813 | 2751 | 210.711 | 424 | 72 | 101 | > 1000 | 236 | 3785
| Mean |  | 1184.5 | 34.6667 | 39.1667 | 186.6667 | 140 | 861.6667 | 678.3333 | 306.6667 | 377 | 169.4128 | 198 | 2770 | 359.2386 | 2694.5 | 202.3028 | 436.3333 | 81.3333 | 100.3333 |  | 234.8333 | 3763.6667
| Median |  | 1225.5 | 33 | 42 | 180 | 105 | 865 | 770 | 300 | 382.5 | 169.9949 | 242 | 2565 | 351.9754 | 2727 | 200.744 | 436.5 | 73.5 | 98.5 |  | 235 | 3768
| Std Dev. |  | 341.4831 | 9.4375 | 5.6362 | 25.8199 | 109.3618 | 264.531 | 259.6472 | 53.1664 | 15.7099 | 4.7038 | 82.5712 | 518.6521 | 39.575 | 88.4347 | 4.4515 | 7.8909 | 34.6506 | 10.5956 |  | 1.7224 | 30.2104
| Rel.Std.Dev. |  | 28.83% | 27.22% | 14.39% | 13.83% | 78.12% | 30.7% | 38.28% | 17.34% | 4.167% | 2.777% | 41.7% | 18.72% | 11.02% | 3.282% | 2.2% | 1.808% | 42.6% | 10.56% |  | 0.7335% | 0.8027%
| PDM3 |  | 350.2% | -86.82% | -85.11% | -29.05% | -46.79% | 227.5% | 157.8% | 16.56% | 43.3% | -35.61% | -24.74% | 952.9% | 36.55% | 924.2% | -23.1% | 65.85% | -69.09% | -61.86% |  | -10.74% | 1331%
| Z-Score (Absolute) |  | 3.67 | 0.95 | 0.93 | 0.34 | 0.52 | 2.37 | 1.64 | 0.14 | 0.43 | 0.41 | 0.29 | 10.03 | 0.36 | 9.73 | 0.27 | 0.66 | 0.76 | 0.68 | NA | 0.14 | 14.02
| Analytical results for Be in OREAS 611 (Certified Value 1.46 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 1.6 | 2 | 1.6 | 1.39 | 1.48 | 1.5 | 1.47 | 1.44 | 1.5 | 1.33 | 1.4242 | 2 | 1.95 | 1.7 | 1.2 | 1.295 | 1.5 | 1.5 | 1.4 | 2 | < 0.5 |  |  | 1
| 1 | 2 | 1.6 | 1 | 1.6 | 1.38 | 1.54 | 1.4 | 1.35 | 1.42 | 1.6 | 1.42 | 1.4499 | 2 | 1.98 | 1.4 | 1 | 1.251 | 1.5 | 1.5 | 1.3 | 1 | < 0.5 |  |  | 32
| 1 | 3 | 1.7 | 1 | 1.5 | 1.27 | 1.39 | 1.53 | 1.37 | 1.5 | 1.5 | 1.4 | 1.4715 | 2 | 1.88 | 1.5 | 2 | 1.274 | 1.4 | 1.5 | 1.4 | 2 | < 0.5 |  |  | 16
| 1 | 4 | 1.6 | 2 | 1.6 | 1.48 | 1.44 | 1.42 | 1.42 | 1.39 | 1.5 | 1.37 | 1.4516 | 2 | 1.94 | 1.7 | 1.2 | 1.273 | 1.4 | 1.5 | 1.4 | 2 | < 0.5 |  |  | 1.4613
| 1 | 5 | 1.7 | < 1 | 1.5 | 1.68 | 1.56 | 1.45 | 1.53 | 1.47 | 1.6 | 1.32 | 1.4527 | 2 | 1.91 | 1.5 | 1.1 | 1.284 | 1.5 | 1.4 | 1.4 | 2 | < 0.5 |  |  | 16
| 1 | 6 | 1.7 | 1 | 1.6 | 1.34 | 1.6 | 1.45 | 1.39 | 1.48 | 1.6 | 1.44 | 1.437 | 2 | 1.91 | 1.5 | 1.1 | 1.252 | 1.4 | 1.4 | 1.4 | 2 | < 0.5
| Mean |  | 1.65 | 1.4 | 1.5667 | 1.4233 | 1.5017 | 1.4583 | 1.4217 | 1.45 | 1.55 | 1.38 | 1.4478 | 2 | 1.9283 | 1.55 | 1.2667 | 1.2715 | 1.45 | 1.4667 | 1.3833 | 1.8333
| Median |  | 1.65 | 1 | 1.6 | 1.385 | 1.51 | 1.45 | 1.405 | 1.455 | 1.55 | 1.385 | 1.4507 | 2 | 1.925 | 1.5 | 1.15 | 1.2735 | 1.45 | 1.5 | 1.4 | 2
| Std Dev. |  | 0.054772 | 0.54772 | 0.05164 | 0.1432 | 0.079099 | 0.048751 | 0.067651 | 0.040988 | 0.054772 | 0.04858 | 0.015994 | 0 | 0.035449 | 0.12247 | 0.36697 | 0.017421 | 0.054772 | 0.05164 | 0.040825 | 0.40825
| Rel.Std.Dev. |  | 3.32% | 39.12% | 3.296% | 10.06% | 5.267% | 3.343% | 4.759% | 2.827% | 3.534% | 3.52% | 1.105% | 0% | 1.838% | 7.902% | 28.97% | 1.37% | 3.777% | 3.521% | 2.951% | 22.27%
| PDM3 |  | 12.91% | -4.195% | 7.21% | -2.599% | 2.762% | -0.2035% | -2.713% | -0.7738% | 6.069% | -5.564% | -0.9256% | 36.86% | 31.96% | 6.069% | -13.32% | -12.99% | -0.7738% | 0.3667% | -5.336% | 25.46%
| Z-Score (Absolute) |  | 1.97 | NA | 1.15 | 0.26 | 0.51 | 0.08 | 0.28 | 0 | 0.99 | 0.69 | 0.02 | NA | 4.72 | 0.99 | 1.81 | 1.76 | 0 | 0.16 | 0.66 | NA | 11.84
Results from laboratories 3, 14 & 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 611 (Certified Value 265 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 241.26 | 270.02 | 222 | 265 | 265 | 276 | 264 | 271 | 263 | 264.75 | 288.9679 | 263.78 | 255 | 268 | 262.7358 | 257.96 | 267.171 | 256.3 | 261 | 252 | 261.1 | 240 | 169.5 | 1
| 1 | 2 | 237.86 | 285.13 | 227 | 263 | 258 | 272 | 258 | 277 | 242 | 272.05 | 280.2759 | 281.14 | 260 | 271 | 261.5219 | 247.33 | 267.5 | 256.6 | 261 | 244 | 265.7 | 240 | 173.2 | 33
| 1 | 3 | 245.41 | 282.19 | 223 | 273 | 261 | 268 | 264 | 271 | 263 | 273.06 | 290.2125 | 287.71 | 256 | 268 | 262.5387 | 256.71 | 269.091 | 251.2 | 258 | 250 | 265.9 | 240 | 173.9 | 16
| 1 | 4 | 248.48 | 293.47 | 226 | 266 | 267 | 264 | 270 | 275 | 262 | 266.73 | 278.1475 | 283.02 | 259 | 268 | 262.8135 | 257.8 | 268.17 | 256.1 | 261 | 257 | 261.8 | 240 | 170.6 | 265.4204
| 1 | 5 | 244.06 | 281.98 | 230 | 282 | 262 | 263 | 263 | 285 | 267 | 270.35 | 289.474 | 284.1 | 268 | 268 | 266.9903 | 270.07 | 263.094 | 251.9 | 252 | 248 | 260.5 | 230 | 174.3 | 17
| 1 | 6 | 248.74 | 262.28 | 235 | 273 | 263 | 269 | 260 | 276 | 256 | 271.06 | 281.0583 | 282.29 | 263 | 266 | 262.7147 | 265.7 | 267.408 | 251.1 | 261 | 250 | 255.4 | 230 | 172
| Mean |  | 244.3017 | 279.1783 | 227.1667 | 270.3333 | 262.6667 | 268.6667 | 263.1667 | 275.8333 | 258.8333 | 269.6667 | 284.6893 | 280.34 | 260.1667 | 268.1667 | 263.2191 | 259.2617 | 267.0723 | 253.8667 | 259 | 250.1667 | 261.7333 | 236.6667 | 172.25
| Median |  | 244.735 | 282.085 | 226.5 | 269.5 | 262.5 | 268.5 | 263.5 | 275.5 | 262.5 | 270.705 | 285.0131 | 282.655 | 259.5 | 268 | 262.7252 | 257.88 | 267.454 | 254 | 261 | 250 | 261.45 | 240 | 172.6
| Std Dev. |  | 4.2235 | 11.1895 | 4.7924 | 7.0899 | 3.1411 | 4.8854 | 4.1191 | 5.1543 | 8.9759 | 3.2373 | 5.4252 | 8.418 | 4.7924 | 1.6021 | 1.9088 | 7.8864 | 2.0694 | 2.7208 | 3.6332 | 4.3089 | 3.8764 | 5.164 | 1.9066
| Rel.Std.Dev. |  | 1.729% | 4.008% | 2.11% | 2.623% | 1.196% | 1.818% | 1.565% | 1.869% | 3.468% | 1.2% | 1.906% | 3.003% | 1.842% | 0.5974% | 0.7252% | 3.042% | 0.7749% | 1.072% | 1.403% | 1.722% | 1.481% | 2.182% | 1.107%
| PDM3 |  | -7.957% | 5.183% | -14.41% | 1.851% | -1.037% | 1.223% | -0.8491% | 3.923% | -2.482% | 1.6% | 7.26% | 5.621% | -1.979% | 1.035% | -0.8293% | -2.32% | 0.6224% | -4.353% | -2.419% | -5.747% | -1.389% | -10.83% | -35.1%
| Z-Score (Absolute) |  | 1.77 | 1.59 | 3.42 | 0.74 | 0 | 0.58 | 0.05 | 1.27 | 0.37 | 0.67 | 2.12 | 1.7 | 0.24 | 0.53 | 0.05 | 0.33 | 0.42 | 0.85 | 0.35 | 1.2 | 0.09 | 2.5 | 8.71
| Analytical results for Ca in OREAS 611 (Certified Value 0.227 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.2301 | 0.22 | 0.2 | 0.22 | 0.23 | 0.24 | 0.22 | 0.23 | 0.248 | 0.2352 | 0.24678 | 0.1801 | 0.24 | 0.23 | 0.23651 | 0.17 | 0.171 | 0.22 | 0.23 | 0.21 | 0.24 | 0.13 | 0.2468 | 1
| 1 | 2 | 0.2248 | 0.21 | 0.2 | 0.23 | 0.23 | 0.24 | 0.22 | 0.24 | 0.251 | 0.2419 | 0.24319 | 0.1955 | 0.24 | 0.23 | 0.23436 | 0.18 | 0.171 | 0.21 | 0.23 | 0.2 | 0.24 | 0.13 | 0.2617 | #N/A
| 1 | 3 | 0.2277 | 0.21 | 0.2 | 0.23 | 0.22 | 0.23 | 0.22 | 0.23 | 0.258 | 0.2429 | 0.25327 | 0.2001 | 0.24 | 0.24 | 0.23009 | 0.22 | 0.17 | 0.2 | 0.23 | 0.21 | 0.24 | 0.14 | 0.2666 | 16
| 1 | 4 | 0.224 | 0.21 | 0.21 | 0.23 | 0.23 | 0.24 | 0.23 | 0.24 | 0.236 | 0.2324 | 0.23467 | 0.1838 | 0.24 | 0.23 | 0.22844 | 0.19 | 0.179 | 0.2 | 0.23 | 0.21 | 0.23 | 0.14 | 0.2815 | 0.22729
| 1 | 5 | 0.2281 | 0.22 | 0.22 | 0.24 | 0.23 | 0.23 | 0.22 | 0.25 | 0.234 | 0.2393 | 0.23884 | 0.1856 | 0.24 | 0.23 | 0.23428 | 0.22 | 0.161 | 0.21 | 0.22 | 0.2 | 0.23 | 0.13 | 0.25306 | 18
| 1 | 6 | 0.2299 | 0.2 | 0.23 | 0.23 | 0.23 | 0.23 | 0.22 | 0.24 | 0.249 | 0.2349 | 0.24879 | 0.2048 | 0.24 | 0.23 | 0.23621 | 0.21 | 0.167 | 0.21 | 0.23 | 0.22 | 0.23 | 0.13 | 0.26213
| Mean |  | 0.22743 | 0.21167 | 0.21 | 0.23 | 0.22833 | 0.235 | 0.22167 | 0.23833 | 0.246 | 0.23777 | 0.24426 | 0.19165 | 0.24 | 0.23167 | 0.23332 | 0.19833 | 0.16983 | 0.20833 | 0.22833 | 0.20833 | 0.235 | 0.13333 | 0.26197
| Median |  | 0.2279 | 0.21 | 0.205 | 0.23 | 0.23 | 0.235 | 0.22 | 0.24 | 0.2485 | 0.23725 | 0.24499 | 0.19055 | 0.24 | 0.23 | 0.23432 | 0.2 | 0.1705 | 0.21 | 0.23 | 0.21 | 0.235 | 0.13 | 0.26192
| Std Dev. |  | 0.0025469 | 0.0075277 | 0.012649 | 0.0063246 | 0.0040825 | 0.0054772 | 0.0040825 | 0.0075277 | 0.0092304 | 0.0042283 | 0.0067881 | 0.0099073 | 0 | 0.0040825 | 0.0033116 | 0.02137 | 0.0058793 | 0.0075277 | 0.0040825 | 0.0075277 | 0.0054772 | 0.005164 | 0.011937
| Rel.Std.Dev. |  | 1.12% | 3.556% | 6.023% | 2.75% | 1.788% | 2.331% | 1.842% | 3.158% | 3.752% | 1.778% | 2.779% | 5.169% | 0% | 1.762% | 1.419% | 10.77% | 3.462% | 3.613% | 1.788% | 3.613% | 2.331% | 3.873% | 4.557%
| PDM3 |  | 0.06379% | -6.873% | -7.606% | 1.193% | 0.4598% | 3.393% | -2.473% | 4.859% | 8.233% | 4.61% | 7.466% | -15.68% | 5.593% | 1.926% | 2.652% | -12.74% | -25.28% | -8.34% | 0.4598% | -8.34% | 3.393% | -41.34% | 15.26%
| Z-Score (Absolute) |  | 0.05 | 0.96 | 1.06 | 0.1 | 0 | 0.39 | 0.39 | 0.58 | 1.02 | 0.55 | 0.92 | 2.12 | 0.67 | 0.19 | 0.29 | 1.73 | 3.38 | 1.16 | 0 | 1.16 | 0.39 | 5.49 | 1.94
| Analytical results for Cd in OREAS 611 (Certified Value 13.9 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 12.48 | 14.95 | 13 | 13.7 | 13.75 | 13.65 | 13.85 | 14.2 | 13.6 | 14.36 | 13.9146 | 15 | 14.3 | 14.5 | 14.1169 | 11.41 | 14.481 | 14.9 | 12.37 | 13 | 11.66 | 36 | 15.08 | 1
| 1 | 2 | 12.59 | 15.16 | 13.2 | 13.7 | 13.55 | 13.5 | 13.35 | 14.3 | 14 | 13.97 | 13.1537 | 13.49 | 14.1 | 14.5 | 14.2311 | 11.6 | 14.459 | 15 | 12.35 | 13 | 11.81 | 36 | 14.85 | 34
| 1 | 3 | 13.01 | 14.8 | 13.1 | 13.35 | 13.4 | 13.55 | 13.55 | 14.25 | 13.2 | 14.56 | 13.4875 | 14.19 | 14 | 14 | 14.0866 | 11.92 | 14.558 | 14.7 | 12.36 | 13 | 11.86 | 36 | 15.62 | 16
| 1 | 4 | 12.88 | 14.77 | 13.2 | 13.75 | 13.6 | 13.25 | 13.9 | 14.25 | 13.4 | 13.8 | 13.795 | 14.2 | 14.4 | 14.5 | 14.0934 | 11.73 | 14.237 | 14.8 | 12.07 | 14 | 11.88 | 36 | 14.39 | 13.912
| 1 | 5 | 13.08 | 14.43 | 13.9 | 14.05 | 13.95 | 13.65 | 13.95 | 14.8 | 14 | 14.27 | 12.8768 | 13.29 | 14.2 | 14 | 14.3216 | 12.3 | 14.505 | 14.2 | 11.95 | 13 | 11.59 | 35 | 14.86 | 19
| 1 | 6 | 12.8 | 14.16 | 13.3 | 13.95 | 13.7 | 13.75 | 13.7 | 14.5 | 14.1 | 14.26 | 13.5231 | 15.27 | 14.1 | 14 | 14.3788 | 12.04 | 14.191 | 14.6 | 12.3 | 13 | 11.91 | 35 | 15.63
| Mean |  | 12.8067 | 14.7117 | 13.2833 | 13.75 | 13.6583 | 13.5583 | 13.7167 | 14.3833 | 13.7167 | 14.2033 | 13.4585 | 14.24 | 14.1833 | 14.25 | 14.2047 | 11.8333 | 14.4052 | 14.7 | 12.2333 | 13.1667 | 11.785 | 35.6667 | 15.0717
| Median |  | 12.84 | 14.785 | 13.2 | 13.725 | 13.65 | 13.6 | 13.775 | 14.275 | 13.8 | 14.265 | 13.5053 | 14.195 | 14.15 | 14.25 | 14.174 | 11.825 | 14.47 | 14.75 | 12.325 | 13 | 11.835 | 36 | 14.97
| Std Dev. |  | 0.23458 | 0.3613 | 0.31885 | 0.2429 | 0.18819 | 0.1744 | 0.23166 | 0.22949 | 0.37103 | 0.27442 | 0.38926 | 0.78837 | 0.1472 | 0.27386 | 0.12548 | 0.3201 | 0.15239 | 0.28284 | 0.17874 | 0.40825 | 0.13004 | 0.5164 | 0.48396
| Rel.Std.Dev. |  | 1.832% | 2.456% | 2.4% | 1.767% | 1.378% | 1.286% | 1.689% | 1.596% | 2.705% | 1.932% | 2.892% | 5.536% | 1.038% | 1.922% | 0.8834% | 2.705% | 1.058% | 1.924% | 1.461% | 3.101% | 1.103% | 1.448% | 3.211%
| PDM3 |  | -7.945% | 5.748% | -4.519% | -1.165% | -1.824% | -2.542% | -1.404% | 3.388% | -1.404% | 2.094% | -3.26% | 2.358% | 1.95% | 2.429% | 2.104% | -14.94% | 3.545% | 5.664% | -12.07% | -5.358% | -15.29% | 156.4% | 8.336%
| Z-Score (Absolute) |  | 1.3 | 1.32 | 0.64 | 0 | 0.13 | 0.26 | 0.05 | 0.87 | 0.05 | 0.62 | 0.4 | 0.67 | 0.6 | 0.69 | 0.63 | 2.64 | 0.9 | 1.31 | 2.09 | NA | 2.7 | NA | 1.82
Results from laboratories 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 611 (Certified Value 46.4 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 1
2
| 1 | 1 | 38.63 | 33.4 | 43.8 | 44.1 | 48.9 | 46.1 | 49 | 47.22 | 31.02 | 47.3 | 54.2 | 43.5988 | 41.7 | 35.76 | 48.73 | 48 | [7 empty] | 1
| 1 | 2 | 40.17 | 34.6 | 45 | 44.8 | 48.2 | 43.9 | 50.1 | 49.34 | 32.5 | 51.2 | 54.5 | 43.9199 | 44 | 37.81 | 48.5 | 48 | [7 empty] | 35
| 1 | 3 | 36.2 | 29.2 | 45.7 | 43.8 | 48.2 | 46.1 | 49.6 | 50 | 35.18 | 50 | 52.7 | 44.4452 | 47.8 | 38.03 | 49.11 | 47 | [7 empty] | 16
| 1 | 4 | 40.4 | 28.6 | 43.8 | 44.8 | 48.4 | 50.1 | 50.7 | 48.92 | 33.24 | 53.9 | 53.4 | 43.5524 | 47.5 | 38.45 | 48.07 | 48 | [7 empty] | 46.4045
| 1 | 5 | 39.93 | 30.4 | 50.9 | 44.3 | 45.2 | 47.6 | 54.4 | 49.85 | 32.8 | 53.8 | 52.7 | 45.8244 | 46.1 | 36.98 | 47.01 | 46 | [7 empty] | 20
| 1 | 6 | 37.79 | 31.5 | 46.2 | 46.9 | 46.5 | 48.3 | 51.3 | 49.34 | 33.76 | 50.5 | 51.7 | 45.2693 | 44.4 | 38.51 | 48.48 | 47
| Mean |  | 38.8533 | 31.2833 | 45.9 | 44.7833 | 47.5667 | 47.0167 | 50.85 | 49.1117 | 33.0833 | 51.1167 | 53.2 | 44.435 | 45.25 | 37.59 | 48.3167 | 47.3333
| Median |  | 39.28 | 30.95 | 45.35 | 44.55 | 48.2 | 46.85 | 50.4 | 49.34 | 33.02 | 50.85 | 53.05 | 44.1825 | 45.25 | 37.92 | 48.49 | 47.5
| Std Dev. |  | 1.6435 | 2.36 | 2.6367 | 1.1089 | 1.4152 | 2.1395 | 1.917 | 1.0054 | 1.3825 | 2.4959 | 1.0469 | 0.93491 | 2.3313 | 1.0535 | 0.72508 | 0.8165
| Rel.Std.Dev. |  | 4.23% | 7.544% | 5.744% | 2.476% | 2.975% | 4.551% | 3.77% | 2.047% | 4.179% | 4.883% | 1.968% | 2.104% | 5.152% | 2.803% | 1.501% | 1.725%
| PDM3 |  | -16.27% | -32.59% | -1.087% | -3.494% | 2.504% | 1.319% | 9.58% | 5.834% | -28.71% | 10.15% | 14.64% | -4.244% | -2.488% | -18.99% | 4.121% | 2.002%
| Z-Score (Absolute) |  | 2.19 | 4.38 | 0.16 | 0.48 | 0.32 | 0.16 | 1.27 | 0.77 | 3.86 | 1.34 | 1.94 | 0.58 | 0.35 | 2.56 | 0.54 | 0.25
| Analytical results for Co in OREAS 611 (Certified Value 8.7 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 9 | 8.8 | 7.8 | 8.7 | 8.9 | 8.5 | 8.7 | 8.3 | 9 | 8.9 | 8.4809 | 8.2 | 9.06 | 15 | 8.0368 | 8.6 | 9.084 | 8.6 | 8.9 | 8 | 8.7 | 9 | 4.882 | 1
| 1 | 2 | 9.1 | 8.5 | 8.8 | 8.6 | 8.7 | 8.4 | 8.5 | 8.1 | 9 | 9.2 | 8.4424 | 8.5 | 9.09 | 15 | 8.6865 | 8.72 | 9.18 | 8.8 | 8.5 | 8 | 8.7 | 9 | 4.931 | 21
| 1 | 3 | 9.6 | 8.8 | 8.5 | 8.4 | 8.7 | 8.5 | 8.8 | 8.1 | 9 | 9.1 | 8.6935 | 9 | 9.2 | 15 | 8.3636 | 8.86 | 8.894 | 8.8 | 8.4 | 8 | 9 | 9 | 4.98 | 16
| 1 | 4 | 9.2 | 9.1 | 8.5 | 8.8 | 8.7 | 8.5 | 8.7 | 8.4 | 9 | 9 | 8.7459 | 8.6 | 9.18 | 15 | 8.0805 | 8.54 | 8.506 | 8.6 | 8.5 | 8 | 8.9 | 9 | 5.029 | 8.6962
| 1 | 5 | 9.5 | 9.4 | 8.6 | 9.3 | 8.8 | 8.6 | 8.8 | 8.5 | 9 | 9.2 | 8.8224 | 9 | 8.91 | 15 | 8.0562 | 9.02 | 9.068 | 8.4 | 8.5 | 8 | 8.9 | 9 | 5 | 21
| 1 | 6 | 9.4 | 8 | 8.9 | 8.7 | 8.6 | 8.8 | 8.4 | 8.3 | 9 | 9.4 | 8.4007 | 8.7 | 9.02 | 10 | 8.0812 | 8.86 | 9.034 | 8.6 | 8.6 | 8 | 8.6 | 9 | 4.9127
| Mean |  | 9.3 | 8.7667 | 8.5167 | 8.75 | 8.7333 | 8.55 | 8.65 | 8.2833 | 9 | 9.1333 | 8.5976 | 8.6667 | 9.0767 | 14.1667 | 8.2174 | 8.7667 | 8.961 | 8.6333 | 8.5667 | 8 | 8.8 | 9 | 4.9558
| Median |  | 9.3 | 8.8 | 8.55 | 8.7 | 8.7 | 8.5 | 8.7 | 8.3 | 9 | 9.15 | 8.5872 | 8.65 | 9.075 | 15 | 8.0808 | 8.79 | 9.051 | 8.6 | 8.5 | 8 | 8.8 | 9 | 4.9555
| Std Dev. |  | 0.23664 | 0.48442 | 0.38687 | 0.30166 | 0.10328 | 0.13784 | 0.16432 | 0.16021 | 0 | 0.17512 | 0.17789 | 0.30768 | 0.10708 | 2.0412 | 0.25974 | 0.18052 | 0.24141 | 0.15055 | 0.17512 | 0 | 0.15492 | 0 | 0.056243
| Rel.Std.Dev. |  | 2.545% | 5.526% | 4.542% | 3.448% | 1.183% | 1.612% | 1.9% | 1.934% | 0% | 1.917% | 2.069% | 3.55% | 1.18% | 14.41% | 3.161% | 2.059% | 2.694% | 1.744% | 2.044% | 0% | 1.76% | 0% | 1.135%
| PDM3 |  | 6.943% | 0.8103% | -2.064% | 0.6187% | 0.427% | -1.681% | -0.5312% | -4.748% | 3.494% | 5.027% | -1.133% | -0.3396% | 4.375% | 62.91% | -5.505% | 0.8103% | 3.045% | -0.7229% | -1.49% | -8.006% | 1.194% | 3.494% | -43.01%
| Z-Score (Absolute) |  | 3.66 | 0.58 | 0.87 | 0.48 | 0.39 | 0.67 | 0.1 | 2.22 | NA | 2.7 | 0.4 | 0 | 2.37 | NA | 2.6 | 0.58 | 1.7 | 0.19 | 0.58 | NA | 0.77 | NA | 21.45
Results from laboratories 10, 16, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 611 (Certified Value 51 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 51 | 42 | 43 | 52 | 53 | 51 | 53 | 57 | 49 | 60 | 50.9628 | 58 | 65 | 30 | 48.2834 | 51 | 36.665 | 53 | 41 | 45 | 61 | 53 | 57.77 | 1
| 1 | 2 | 48 | 43 | 42 | 51 | 53 | 50 | 55 | 55 | 49 | 61 | 49.7967 | 62 | 48.4 | 30 | 32.1871 | 53 | 45.661 | 51 | 43 | 45 | 61 | 55 | 58.3 | 37
| 1 | 3 | 47 | 45 | 40 | 54 | 53 | 50 | 54 | 56 | 48 | 63 | 50.4184 | 68 | 46.3 | 40 | 32.2292 | 52 | 36.824 | 50 | 49 | 42 | 50 | 53 | 58.83 | 16
| 1 | 4 | 49 | 47 | 41 | 52 | 54 | 49 | 54 | 55 | 49 | 58 | 51.7957 | 60 | 46.1 | 40 | 33.5954 | 51 | 35.088 | 50 | 37 | 46 | 62 | 54 | 59.36 | 50.5179
| 1 | 5 | 50 | 46 | 39 | 51 | 53 | 53 | 53 | 57 | 50 | 59 | 51.4796 | 68 | 46.7 | 40 | 54.3931 | 53 | 38.664 | 50 | 46 | 48 | 61 | 54 | 59.89 | 22
| 1 | 6 | 51 | 43 | 60 | 51 | 55 | 54 | 51 | 58 | 52 | 61 | 52.0651 | 58 | 48.8 | 40 | 42.7032 | 52 | 35.866 | 50 | 42 | 43 | 57 | 54 | 60.42
| Mean |  | 49.3333 | 44.3333 | 44.1667 | 51.8333 | 53.5 | 51.1667 | 53.3333 | 56.3333 | 49.5 | 60.3333 | 51.0864 | 62.3333 | 50.2167 | 36.6667 | 40.5652 | 52 | 38.128 | 50.6667 | 43 | 44.8333 | 58.6667 | 53.8333 | 59.095
| Median |  | 49.5 | 44 | 41.5 | 51.5 | 53 | 50.5 | 53.5 | 56.5 | 49 | 60.5 | 51.2212 | 61 | 47.55 | 40 | 38.1493 | 52 | 36.7445 | 50 | 42.5 | 45 | 61 | 54 | 59.095
| Std Dev. |  | 1.633 | 1.9664 | 7.8846 | 1.169 | 0.83666 | 1.9408 | 1.3663 | 1.2111 | 1.3784 | 1.7512 | 0.86491 | 4.6332 | 7.3281 | 5.164 | 9.4193 | 0.89443 | 3.879 | 1.2111 | 4.1473 | 2.137 | 4.5898 | 0.75277 | 0.99154
| Rel.Std.Dev. |  | 3.31% | 4.435% | 17.85% | 2.255% | 1.564% | 3.793% | 2.562% | 2.15% | 2.785% | 2.903% | 1.693% | 7.433% | 14.59% | 14.08% | 23.22% | 1.72% | 10.17% | 2.39% | 9.645% | 4.766% | 7.824% | 1.398% | 1.678%
| PDM3 |  | -2.345% | -12.24% | -12.57% | 2.604% | 5.903% | 1.284% | 5.573% | 11.51% | -2.015% | 19.43% | 1.125% | 23.39% | -0.5964% | -27.42% | -19.7% | 2.934% | -24.53% | 0.2944% | -14.88% | -11.25% | 16.13% | 6.563% | 16.98%
| Z-Score (Absolute) |  | 0.31 | 1.16 | 1.19 | 0.12 | 0.4 | 0.01 | 0.38 | 0.89 | 0.28 | 1.57 | 0.01 | 1.91 | 0.16 | NA | 1.8 | 0.15 | 2.22 | 0.08 | 1.38 | 1.07 | 1.28 | 0.46 | 1.36
Results from laboratory 16 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 611 (Certified Value 2.06 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.9 | 2.1 | 2.04 | 2.2 | 1.97 | 2.06 | 2.12 | 2.05 | 2.1 | 2.09 | 2.2672 | 1.9 | 1.88 | 2.2 | 1.76 | 1.7 | 2 | 2.2 | [5 empty] | 1
| 1 | 2 | 1.8 | 2.2 | 1.97 | 2.19 | 1.92 | 2.09 | 2.05 | 2.07 | 2.2 | 2.18 | 2.2765 | 1.9 | 1.93 | 2.2 | 1.8 | 1.75 | 2 | 2.2 | [5 empty] | 38
| 1 | 3 | 1.9 | 2.1 | 1.97 | 2.15 | 1.9 | 2.03 | 2.09 | 2.03 | 2.1 | 2.16 | 2.4272 | 2.2 | 1.89 | 2.2 | 1.98 | 1.78 | 2 | 2.2 | [5 empty] | 16
| 1 | 4 | 1.9 | 2.1 | 2.02 | 2.15 | 1.99 | 2.04 | 2.22 | 2.1 | 2.2 | 2.1 | 2.2063 | 2.1 | 1.95 | 2.2 | 1.94 | 1.79 | 2 | 2.2 | [5 empty] | 2.0619
| 1 | 5 | 1.9 | 2.1 | 2.09 | 2.27 | 1.98 | 2.01 | 2.13 | 2.12 | 2.1 | 2.14 | 2.3035 | 1.9 | 1.96 | 2.2 | 2.04 | 1.75 | 2 | 2.2 | [5 empty] | 23
| 1 | 6 | 1.9 | 2 | 2.08 | 2.16 | 2.08 | 2.06 | 2.09 | 2.13 | 2.1 | 2.16 | 2.3787 | 2.3 | 1.94 | 2.2 | 1.97 | 1.79 | 2 | 2.2
| Mean |  | 1.8833 | 2.1 | 2.0283 | 2.1867 | 1.9733 | 2.0483 | 2.1167 | 2.0833 | 2.1333 | 2.1383 | 2.3099 | 2.05 | 1.925 | 2.2 | 1.915 | 1.76 | 2 | 2.2
| Median |  | 1.9 | 2.1 | 2.03 | 2.175 | 1.975 | 2.05 | 2.105 | 2.085 | 2.1 | 2.15 | 2.29 | 2 | 1.935 | 2.2 | 1.955 | 1.765 | 2 | 2.2
| Std Dev. |  | 0.040825 | 0.063246 | 0.051929 | 0.045898 | 0.06314 | 0.027869 | 0.05785 | 0.039833 | 0.05164 | 0.036009 | 0.08026 | 0.17607 | 0.032711 | 4.8648e-16 | 0.11023 | 0.034641 | 0 | 4.8648e-16
| Rel.Std.Dev. |  | 2.168% | 3.012% | 2.56% | 2.099% | 3.2% | 1.361% | 2.733% | 1.912% | 2.421% | 1.684% | 3.475% | 8.589% | 1.699% | 2.211e-14% | 5.756% | 1.968% | 0% | 2.211e-14%
| PDM3 |  | -8.658% | 1.85% | -1.626% | 6.053% | -4.293% | -0.6559% | 2.658% | 1.042% | 3.467% | 3.709% | 12.03% | -0.575% | -6.638% | 6.7% | -7.123% | -14.64% | -3% | 6.7%
| Z-Score (Absolute) |  | 1.31 | 0.11 | 0.36 | 0.67 | 0.72 | 0.23 | 0.22 | 0 | 0.33 | 0.36 | 1.48 | 0.22 | 1.03 | 0.76 | 1.1 | 2.11 | NA | 0.76
Results from laboratory 21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 611 (Certified Value 1.17 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*MS | 4A*OES | 4A*AAS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*AAS | 2
3
| 1 | 1 | 1.1816 | 1.151 | 1.18 | 1.12 | 1.13 | 1.18 | 1.18 | 1.175 | 1.14 | 1.1715 | 1.1656 | 1.183 | 1.21 | 1.14 | 1.171 | 1.16 | 1.19 | 1.166 | 1.2 | > 1 | 1.14 | 1.33 | 1.199 | 1
| 1 | 2 | 1.1471 | 1.149 | 1.15 | 1.125 | 1.145 | 1.19 | 1.19 | 1.17 | 1.16 | 1.2027 | 1.1848 | 1.178 | 1.21 | 1.15 | 1.1821 | 1.159 | 1.18 | 1.169 | 1.21 | > 1 | 1.15 | 1.32 | 1.183 | 23
| 1 | 3 | 1.1858 | 1.158 | 1.21 | 1.15 | 1.155 | 1.185 | 1.18 | 1.14 | 1.13 | 1.2058 | 1.1828 | 1.162 | 1.21 | 1.16 | 1.1771 | 1.166 | 1.19 | 1.176 | 1.19 | > 1 | 1.14 | 1.33 | 1.117 | 16
| 1 | 4 | 1.143 | 1.131 | 1.16 | 1.115 | 1.155 | 1.155 | 1.175 | 1.165 | 1.16 | 1.1751 | 1.1755 | 1.158 | 1.19 | 1.15 | 1.1686 | 1.156 | 1.18 | 1.173 | 1.19 | > 1 | 1.14 | 1.32 | 1.17 | 1.1667
| 1 | 5 | 1.1593 | 1.149 | 1.19 | 1.12 | 1.15 | 1.18 | 1.17 | 1.165 | 1.18 | 1.1884 | 1.1641 | 1.184 | 1.18 | 1.14 | 1.1736 | 1.152 | 1.17 | 1.161 | 1.17 | > 1 | 1.14 | 1.31 | 1.1755 | 24
| 1 | 6 | 1.1603 | 1.159 | 1.16 | 1.115 | 1.135 | 1.185 | 1.17 | 1.145 | 1.15 | 1.1956 | 1.1784 | 1.19 | 1.16 | 1.14 | 1.184 | 1.157 | 1.18 | 1.156 | 1.19 | > 1 | 1.13 | 1.28 | 1.1677
| Mean |  | 1.1628 | 1.1495 | 1.175 | 1.1242 | 1.145 | 1.1792 | 1.1775 | 1.16 | 1.1533 | 1.1899 | 1.1752 | 1.1758 | 1.1933 | 1.1467 | 1.1761 | 1.1583 | 1.1817 | 1.1668 | 1.1917 |  | 1.14 | 1.315 | 1.1687
| Median |  | 1.1598 | 1.15 | 1.17 | 1.12 | 1.1475 | 1.1825 | 1.1775 | 1.165 | 1.155 | 1.192 | 1.177 | 1.1805 | 1.2 | 1.145 | 1.1754 | 1.158 | 1.18 | 1.1675 | 1.19 |  | 1.14 | 1.32 | 1.1727
| Std Dev. |  | 0.017547 | 0.010075 | 0.022583 | 0.013197 | 0.010488 | 0.012416 | 0.0075829 | 0.014142 | 0.017512 | 0.014212 | 0.0086612 | 0.012906 | 0.020656 | 0.008165 | 0.0061347 | 0.0046762 | 0.0075277 | 0.0074677 | 0.013292 |  | 0.0063246 | 0.018708 | 0.027729
| Rel.Std.Dev. |  | 1.509% | 0.8764% | 1.922% | 1.174% | 0.916% | 1.053% | 0.644% | 1.219% | 1.518% | 1.194% | 0.737% | 1.098% | 1.731% | 0.7121% | 0.5216% | 0.4037% | 0.637% | 0.64% | 1.115% |  | 0.5548% | 1.423% | 2.373%
| PDM3 |  | -0.3313% | -1.476% | 0.7101% | -3.647% | -1.861% | 1.067% | 0.9244% | -0.5756% | -1.147% | 1.984% | 0.7282% | 0.7815% | 2.281% | -1.718% | 0.8006% | -0.7184% | 1.282% | 0.01013% | 2.139% |  | -2.29% | 12.71% | 0.1701%
| Z-Score (Absolute) |  | 0.48 | 1.19 | 0.17 | 2.54 | 1.43 | 0.39 | 0.3 | 0.63 | 0.98 | 0.96 | 0.18 | 0.21 | 1.14 | 1.34 | 0.22 | 0.72 | 0.52 | 0.27 | 1.05 | NA | 1.69 | 7.61 | 0.17
| Analytical results for Dy in OREAS 611 (Certified Value 1.55 ppm) | [24 empty] | Certified
Dy | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.02: 1.4; 2.03: 1.4; 2.04: 1.64; 2.05: 1.6; 2.06: 1.55 | 1
| 1 | 2 | 2.02: 1.3; 2.03: 1.5; 2.04: 1.72; 2.05: 1.6; 2.06: 1.6 | 10
| 1 | 3 | 2.02: 1.6; 2.03: 1.5; 2.04: 1.72; 2.05: 1.6; 2.06: 1.55 | 16
| 1 | 4 | 2.02: 1.4; 2.03: 1.6; 2.04: 1.64; 2.05: 1.6; 2.06: 1.55 | 1.5513
| 1 | 5 | 2.02: 1.4; 2.03: 1.6; 2.04: 1.65; 2.05: 1.6; 2.06: 1.55 | 25
| 1 | 6 | 2.02: 1.3; 2.03: 1.6; 2.04: 1.67; 2.05: 1.6; 2.06: 1.5
| Mean |  | 2.02: 1.4; 2.03: 1.5333; 2.04: 1.6733; 2.05: 1.6; 2.06: 1.55
| Median |  | 2.02: 1.4; 2.03: 1.55; 2.04: 1.66; 2.05: 1.6; 2.06: 1.55
| Std Dev. |  | 2.02: 0.10954; 2.03: 0.08165; 2.04: 0.037771; 2.05: 2.4324e-16; 2.06: 0.031623
| Rel.Std.Dev. |  | 2.02: 7.825%; 2.03: 5.325%; 2.04: 2.257%; 2.05: 1.52e-14%; 2.06: 2.04%
| PDM3 |  | 2.02: -9.755%; 2.03: -1.16%; 2.04: 7.864%; 2.05: 3.137%; 2.06: -0.08595%
| Z-Score (Absolute) |  | 2.02: 2.02; 2.03: 0.22; 2.04: 1.66; 2.05: 0.67; 2.06: 0
| Analytical results for Er in OREAS 611 (Certified Value 0.61 ppm) | [24 empty] | Certified
Er | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.02: 0.7; 2.03: 0.6; 2.04: 0.62; 2.05: 0.6; 2.06: 0.6 | 1
| 1 | 2 | 2.02: 0.7; 2.03: 0.7; 2.04: 0.61; 2.05: 0.5; 2.06: 0.65 | 11
| 1 | 3 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.6; 2.05: 0.6; 2.06: 0.6 | 16
| 1 | 4 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.59; 2.05: 0.6; 2.06: 0.6 | 0.613
| 1 | 5 | 2.02: 0.6; 2.03: 0.7; 2.04: 0.65; 2.05: 0.6; 2.06: 0.6 | 26
| 1 | 6 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.62; 2.05: 0.6; 2.06: 0.6
| Mean |  | 2.02: 0.63333; 2.03: 0.63333; 2.04: 0.615; 2.05: 0.58333; 2.06: 0.60833
| Median |  | 2.02: 0.6; 2.03: 0.6; 2.04: 0.615; 2.05: 0.6; 2.06: 0.6
| Std Dev. |  | 2.02: 0.05164; 2.03: 0.05164; 2.04: 0.020736; 2.05: 0.040825; 2.06: 0.020412
| Rel.Std.Dev. |  | 2.02: 8.154%; 2.03: 8.154%; 2.04: 3.372%; 2.05: 6.999%; 2.06: 3.355%
| PDM3 |  | 2.02: 3.317%; 2.03: 3.317%; 2.04: 0.3263%; 2.05: -4.84%; 2.06: -0.7613%
| Z-Score (Absolute) |  | 2.02: 0.67; 2.03: 0.67; 2.04: 0; 2.05: 1.16; 2.06: 0.25
| Analytical results for Eu in OREAS 611 (Certified Value 0.76 ppm) | [24 empty] | Certified
Eu | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
3
| 1 | 1 | 2.02: 0.7; 2.03: 0.7; 2.04: 0.83; 2.05: 0.8; 2.06: 0.75 | 1
| 1 | 2 | 2.02: 1; 2.03: 0.7; 2.04: 0.82; 2.05: 0.7; 2.06: 0.75 | 12
| 1 | 3 | 2.02: 0.8; 2.03: 0.72; 2.04: 0.82; 2.05: 0.7; 2.06: 0.7 | 16
| 1 | 4 | 2.02: 0.8; 2.03: 0.75; 2.04: 0.83; 2.05: 0.8; 2.06: 0.75 | 0.75567
| 1 | 5 | 2.02: 0.8; 2.03: 0.71; 2.04: 0.79; 2.05: 0.7; 2.06: 0.75 | 27
| 1 | 6 | 2.02: 0.6; 2.03: 0.75; 2.04: 0.81; 2.05: 0.8; 2.06: 0.75
| Mean |  | 2.02: 0.78333; 2.03: 0.72167; 2.04: 0.81667; 2.05: 0.75; 2.06: 0.74167
| Median |  | 2.02: 0.8; 2.03: 0.715; 2.04: 0.82; 2.05: 0.75; 2.06: 0.75
| Std Dev. |  | 2.02: 0.13292; 2.03: 0.023166; 2.04: 0.015055; 2.05: 0.054772; 2.06: 0.020412
| Rel.Std.Dev. |  | 2.02: 16.97%; 2.03: 3.21%; 2.04: 1.844%; 2.05: 7.303%; 2.06: 2.752%
| PDM3 |  | 2.02: 3.661%; 2.03: -4.499%; 2.04: 8.072%; 2.05: -0.7499%; 2.06: -1.853%
| Z-Score (Absolute) |  | 2.02: 0.79; 2.03: 0.67; 2.04: 1.59; 2.05: 0; 2.06: 0.2
| Analytical results for Fe in OREAS 611 (Certified Value 2.54 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.58 | 2.43 | 2.35 | 2.4 | 2.5 | 2.56 | 2.4 | 2.52 | 2.46 | 2.64 | 2.7104 | 2.4475 | 2.57 | 2.71 | 2.4835 | 2.45 | 2.475 | 2.5 | 2.63 | 2.6 | 2.58 | 2.5 | 2.799 | 1
| 1 | 2 | 2.5 | 2.42 | 2.43 | 2.42 | 2.46 | 2.51 | 2.37 | 2.55 | 2.63 | 2.7 | 2.6838 | 2.5208 | 2.59 | 2.74 | 2.4523 | 2.49 | 2.443 | 2.46 | 2.61 | 2.56 | 2.57 | 2.57 | 2.819 | #N/A
| 1 | 3 | 2.54 | 2.43 | 2.43 | 2.47 | 2.46 | 2.49 | 2.43 | 2.52 | 2.44 | 2.74 | 2.739 | 2.5373 | 2.54 | 2.78 | 2.4842 | 2.53 | 2.472 | 2.46 | 2.63 | 2.6 | 2.57 | 2.54 | 2.7895 | 16
| 1 | 4 | 2.5 | 2.43 | 2.5 | 2.43 | 2.51 | 2.5 | 2.44 | 2.53 | 2.57 | 2.62 | 2.6676 | 2.4824 | 2.58 | 2.75 | 2.4745 | 2.45 | 2.461 | 2.48 | 2.65 | 2.59 | 2.53 | 2.55 | 2.7971 | 2.5366
| 1 | 5 | 2.54 | 2.46 | 2.57 | 2.51 | 2.47 | 2.47 | 2.4 | 2.63 | 2.63 | 2.69 | 2.7101 | 2.5099 | 2.55 | 2.7 | 2.4863 | 2.54 | 2.478 | 2.45 | 2.65 | 2.53 | 2.55 | 2.59 | 2.907 | 28
| 1 | 6 | 2.56 | 2.3 | 2.55 | 2.43 | 2.45 | 2.52 | 2.36 | 2.57 | 2.57 | 2.65 | 2.6726 | 2.6237 | 2.54 | 2.72 | 2.4997 | 2.52 | 2.468 | 2.49 | 2.67 | 2.64 | 2.54 | 2.54 | 2.9123
| Mean |  | 2.5367 | 2.4117 | 2.4717 | 2.4433 | 2.475 | 2.5083 | 2.4 | 2.5533 | 2.55 | 2.6733 | 2.6973 | 2.5203 | 2.5617 | 2.7333 | 2.4801 | 2.4967 | 2.4662 | 2.4733 | 2.64 | 2.5867 | 2.5567 | 2.5483 | 2.8373
| Median |  | 2.54 | 2.43 | 2.465 | 2.43 | 2.465 | 2.505 | 2.4 | 2.54 | 2.57 | 2.67 | 2.697 | 2.5154 | 2.56 | 2.73 | 2.4839 | 2.505 | 2.47 | 2.47 | 2.64 | 2.595 | 2.56 | 2.545 | 2.809
| Std Dev. |  | 0.032042 | 0.056362 | 0.083526 | 0.039833 | 0.02429 | 0.030605 | 0.031623 | 0.042269 | 0.082219 | 0.044572 | 0.027394 | 0.059708 | 0.02137 | 0.029439 | 0.015858 | 0.039833 | 0.012797 | 0.019664 | 0.020976 | 0.037771 | 0.019664 | 0.030605 | 0.056896
| Rel.Std.Dev. |  | 1.263% | 2.337% | 3.379% | 1.63% | 0.9814% | 1.22% | 1.318% | 1.655% | 3.224% | 1.667% | 1.016% | 2.369% | 0.8342% | 1.077% | 0.6394% | 1.595% | 0.5189% | 0.795% | 0.7946% | 1.46% | 0.7691% | 1.201% | 2.005%
| PDM3 |  | 0.0009776% | -4.927% | -2.561% | -3.678% | -2.43% | -1.116% | -5.387% | 0.658% | 0.5266% | 5.389% | 6.332% | -0.6455% | 0.9865% | 7.754% | -2.229% | -1.576% | -2.778% | -2.496% | 4.075% | 1.972% | 0.7894% | 0.4609% | 11.85%
| Z-Score (Absolute) |  | 0 | 1.37 | 0.71 | 1.02 | 0.67 | 0.31 | 1.49 | 0.18 | 0.15 | 1.49 | 1.76 | 0.18 | 0.27 | 2.15 | 0.62 | 0.44 | 0.77 | 0.69 | 1.13 | 0.55 | 0.22 | 0.13 | 3.29
| Analytical results for Ga in OREAS 611 (Certified Value 24.4 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 26.4 | 23.43 | 24.4 | 24.9 | 24.7 | 24.9 | 24.5 | 27.4 | 25 | 24.32 | 25.2844 | 23 | 26.8 | 23.6 | 23.246 | 21.1 | 23 | 23.3 | 25 | 23 |  |  |  | 1
| 1 | 2 | 26.1 | 23.91 | 24.9 | 24 | 24.4 | 24.7 | 23.4 | 27.1 | 26.6 | 24.68 | 25.3338 | 20.34 | 27.3 | 24 | 24.5641 | 21.2 | 22.6 | 22.8 | 26 | 23 |  |  |  | 43
| 1 | 3 | 27 | 25.46 | 25.2 | 23.9 | 24.2 | 24.3 | 24.3 | 26.4 | 24.6 | 25.25 | 26.3855 | 22.85 | 27.1 | 24.2 | 24.4343 | 21.6 | 22.5 | 22.7 | 25 | 23 |  |  |  | 16
| 1 | 4 | 25.5 | 25.07 | 25.7 | 24.6 | 24.6 | 24.5 | 25.3 | 26.7 | 25.4 | 24.07 | 25.2227 | 20.95 | 28 | 24 | 22.8908 | 20.8 | 21.6 | 22.9 | 25 | 23 |  |  |  | 24.439
| 1 | 5 | 27 | 24.87 | 26.3 | 25.6 | 24.6 | 24 | 24.6 | 27.4 | 25.2 | 24.82 | 25.9515 | 21.07 | 27 | 23.8 | 23.414 | 22 | 21.8 | 23 | 25 | 23 |  |  |  | 29
| 1 | 6 | 26.7 | 22.06 | 26.3 | 25.1 | 23.6 | 24.3 | 23.8 | 27.5 | 26 | 24.98 | 24.4375 | 25 | 27.8 | 24.2 | 25.5811 | 21.7 | 21.8 | 23.1 | 24 | 23
| Mean |  | 26.45 | 24.1333 | 25.4667 | 24.6833 | 24.35 | 24.45 | 24.3167 | 27.0833 | 25.4667 | 24.6867 | 25.4359 | 22.2017 | 27.3333 | 23.9667 | 24.0217 | 21.4 | 22.2167 | 22.9667 | 25 | 23
| Median |  | 26.55 | 24.39 | 25.45 | 24.75 | 24.5 | 24.4 | 24.4 | 27.25 | 25.3 | 24.75 | 25.3091 | 21.96 | 27.2 | 24 | 23.9241 | 21.4 | 22.15 | 22.95 | 25 | 23
| Std Dev. |  | 0.58224 | 1.2668 | 0.77115 | 0.65549 | 0.40866 | 0.32094 | 0.66156 | 0.4446 | 0.72296 | 0.43265 | 0.66982 | 1.7436 | 0.47188 | 0.23381 | 1.0145 | 0.44272 | 0.56006 | 0.21602 | 0.63246 | 0
| Rel.Std.Dev. |  | 2.201% | 5.249% | 3.028% | 2.656% | 1.678% | 1.313% | 2.721% | 1.642% | 2.839% | 1.753% | 2.633% | 7.854% | 1.726% | 0.9756% | 4.223% | 2.069% | 2.521% | 0.9406% | 2.53% | 0%
| PDM3 |  | 8.229% | -1.251% | 4.205% | 0.9999% | -0.364% | 0.04516% | -0.5004% | 10.82% | 4.205% | 1.014% | 4.079% | -9.155% | 11.84% | -1.933% | -1.707% | -12.43% | -9.093% | -6.024% | 2.296% | -5.888%
| Z-Score (Absolute) |  | 1.31 | 0.17 | 0.68 | 0.18 | 0.03 | 0.03 | 0.05 | 1.72 | 0.68 | 0.18 | 0.66 | 1.41 | 1.88 | 0.28 | 0.24 | 1.92 | 1.4 | 0.92 | 0.38 | 0.9
| Analytical results for Gd in OREAS 611 (Certified Value 2.92 ppm) | [24 empty] | Certified
Gd | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 |  |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.02: 2.8; 2.03: 2.8; 2.04: 3.05; 2.05: 2.6; 2.06: 3.2 |  |  |  | 1
| 1 | 2 | 2.02: 3; 2.03: 2.7; 2.04: 3.15; 2.05: 2.7; 2.06: 3.2 |  |  |  | 13
| 1 | 3 | 2.02: 2.8; 2.03: 2.9; 2.04: 3.18; 2.05: 2.8; 2.06: 3 |  |  |  | 16
| 1 | 4 | 2.02: 2.7; 2.03: 2.6; 2.04: 2.93; 2.05: 3.1; 2.06: 3.2 |  |  |  | 2.9173
| 1 | 5 | 2.02: 3.1; 2.03: 2.8; 2.04: 3.23; 2.05: 2.4; 2.06: 3.2 |  |  |  | 30
| 1 | 6 | 2.02: 2.8; 2.03: 2.7; 2.04: 3.08; 2.05: 2.6; 2.06: 3.2
| Mean |  | 2.02: 2.8667; 2.03: 2.75; 2.04: 3.1033; 2.05: 2.7; 2.06: 3.1667
| Median |  | 2.02: 2.8; 2.03: 2.75; 2.04: 3.115; 2.05: 2.65; 2.06: 3.2
| Std Dev. |  | 2.02: 0.15055; 2.03: 0.10488; 2.04: 0.10727; 2.05: 0.23664; 2.06: 0.08165
| Rel.Std.Dev. |  | 2.02: 5.252%; 2.03: 3.814%; 2.04: 3.457%; 2.05: 8.765%; 2.06: 2.578%
| PDM3 |  | 2.02: -1.737%; 2.03: -5.736%; 2.04: 6.376%; 2.05: -7.45%; 2.06: 8.547%
| Z-Score (Absolute) |  | 2.02: 0; 2.03: 0.47; 2.04: 0.96; 2.05: 0.67; 2.06: 1.21
| Analytical results for Ge in OREAS 611 (Indicative Value 2.35 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.18 | 2.20 | 2.23 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | 6.3 | 1 | 0.18 | 0.22 | 0.21 | 0.14 | 0.16 | 6.2 | 6.3 | 4.5656 | 0.47 | 0.22 | 0.28 | 0.5 | 8.6 | [8 empty] | 1
| 1 | 2 | 6 | 1.3 | 0.2 | 0.23 | 0.2 | 0.15 | 0.18 | 5.9 | 6.5 | 4.3531 | 0.91 | 0.24 | 0.26 | 0.6 | 8.9 | [8 empty] | 6
| 1 | 3 | 6.2 | 2.2 | 0.18 | 0.22 | 0.18 | 0.15 | 0.16 | 6 | 6.4 | 4.189 | 1.26 | 0.23 | 0.13 | 0.5 | 8.6 | [8 empty] | 16
| 1 | 4 | 5.9 | 2 | 0.18 | 0.2 | 0.22 | 0.15 | 0.16 | 5.9 | 6.2 | 3.9733 | 1.11 | 0.24 | 0.31 | 0.6 | 8.6 | [8 empty] | 2.3539
| 1 | 5 | 5.8 | 1.3 | 0.53 | 0.22 | 0.15 | 0.13 | 0.17 | 6.2 | 6.2 | 3.8611 | 1.11 | 0.23 | 0.3 | 0.5 | 8.5 | [8 empty] | 12
| 1 | 6 | 5.4 | 0.6 | 0.48 | 0.24 | 0.14 | 0.12 | 0.18 | 6.2 | 6.3 | 3.7686 | 0.23 | 0.24 | 0.27 | 0.6 | 8.4
| Mean |  | 5.9333 | 1.4 | 0.29167 | 0.22167 | 0.18333 | 0.14 | 0.16833 | 6.0667 | 6.3167 | 4.1184 | 0.84833 | 0.23333 | 0.25833 | 0.55 | 8.6
| Median |  | 5.95 | 1.3 | 0.19 | 0.22 | 0.19 | 0.145 | 0.165 | 6.1 | 6.3 | 4.0811 | 1.01 | 0.235 | 0.275 | 0.55 | 8.6
| Std Dev. |  | 0.32042 | 0.60332 | 0.16618 | 0.013292 | 0.03266 | 0.012649 | 0.0098319 | 0.15055 | 0.1169 | 0.30651 | 0.40882 | 0.008165 | 0.065549 | 0.054772 | 0.16733
| Rel.Std.Dev. |  | 5.4% | 43.09% | 56.98% | 5.996% | 17.81% | 9.035% | 5.841% | 2.482% | 1.851% | 7.442% | 48.19% | 3.499% | 25.37% | 9.959% | 1.946%
| PDM3 |  | 152.1% | -40.52% | -87.61% | -90.58% | -92.21% | -94.05% | -92.85% | 157.7% | 168.4% | 74.96% | -63.96% | -90.09% | -89.03% | -76.63% | 265.4%
| Z-Score (Absolute) |  | 9.51 | 1.5 | 0.46 | 0.58 | 0.65 | 0.72 | 0.67 | 9.75 | 10.19 | 6.3 | 0.53 | 0.56 | 0.52 | 0 | 14.22
| Analytical results for Hf in OREAS 611 (Certified Value 2.26 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.9 | 2.11 | 2.5 | 2.3 | 2.3 | 2.3 | 2.2 | 2.2 | 2.4 | 2.34 | 2.3762 | 2.27 | 2.3 | 2.6 | 1.97 | 2.4 | 2.27 | 2.2 | [5 empty] | 1
| 1 | 2 | 1.9 | 2.15 | 2.4 | 2.3 | 2.3 | 2.2 | 2.2 | 2.3 | 2.3 | 2.45 | 2.3377 | 2.19 | 2.3 | 2.8 | 2.19 | 2.45 | 2.19 | 2.3 | [5 empty] | 29
| 1 | 3 | 2 | 2.1 | 2.4 | 2.3 | 2.2 | 2.2 | 2.3 | 2.1 | 2.3 | 2.43 | 2.3669 | 2.33 | 2.2 | 2.6 | 2.09 | 2.3 | 2.15 | 2.2 | [5 empty] | 16
| 1 | 4 | 2 | 2.23 | 2.4 | 2.2 | 2.3 | 2.3 | 2.3 | 2.2 | 2.4 | 2.32 | 2.2749 | 2.4 | 2.3 | 2.8 | 2.06 | 2.34 | 2.16 | 2.3 | [5 empty] | 2.2641
| 1 | 5 | 2.1 | 2.03 | 2.5 | 2.3 | 2.2 | 2.1 | 2.2 | 2.2 | 2.4 | 2.33 | 2.3056 | 2.15 | 2.3 | 2.6 | 2.01 | 2.35 | 2.15 | 2.1 | [5 empty] | 31
| 1 | 6 | 2 | 2.13 | 2.6 | 2.3 | 2.4 | 2.1 | 2.3 | 2.2 | 2.3 | 2.42 | 2.228 | 2.4 | 2.2 | 2.6 | 2.13 | 2.4 | 2.18 | 2.1
| Mean |  | 1.9833 | 2.125 | 2.4667 | 2.2833 | 2.2833 | 2.2 | 2.25 | 2.2 | 2.35 | 2.3817 | 2.3149 | 2.29 | 2.2667 | 2.6667 | 2.075 | 2.3733 | 2.1833 | 2.2
| Median |  | 2 | 2.12 | 2.45 | 2.3 | 2.3 | 2.2 | 2.25 | 2.2 | 2.35 | 2.38 | 2.3217 | 2.3 | 2.3 | 2.6 | 2.075 | 2.375 | 2.17 | 2.2
| Std Dev. |  | 0.075277 | 0.065651 | 0.08165 | 0.040825 | 0.075277 | 0.089443 | 0.054772 | 0.063246 | 0.054772 | 0.057764 | 0.056982 | 0.10564 | 0.05164 | 0.10328 | 0.079937 | 0.053541 | 0.045461 | 0.089443
| Rel.Std.Dev. |  | 3.795% | 3.089% | 3.31% | 1.788% | 3.297% | 4.066% | 2.434% | 2.875% | 2.331% | 2.425% | 2.462% | 4.613% | 2.278% | 3.873% | 3.852% | 2.256% | 2.082% | 4.066%
| PDM3 |  | -12.4% | -6.144% | 8.946% | 0.8487% | 0.8487% | -2.832% | -0.6236% | -2.832% | 3.793% | 5.192% | 2.243% | 1.143% | 0.1126% | 17.78% | -8.353% | 4.824% | -3.568% | -2.832%
| Z-Score (Absolute) |  | 2.62 | 1.35 | 1.72 | 0.07 | 0.07 | 0.67 | 0.22 | 0.67 | 0.67 | 0.96 | 0.36 | 0.13 | 0.07 | 3.52 | 1.8 | 0.88 | 0.82 | 0.67
| Analytical results for Hg in OREAS 611 (Indicative Value 0.66 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | 2.02: Lab; 2.03: Lab | [5 empty] | 1
| No | No | 2.02: 2.04; 2.03: 2.17 | [5 empty] | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 2.02: 0.84; 2.03: < 2 | [5 empty] | 1
| 1 | 2 | 2.02: 0.77; 2.03: < 2 | [5 empty] | 5
| 1 | 3 | 2.02: 0.21; 2.03: < 2 | [5 empty] | 16
| 1 | 4 | 2.02: 0.53; 2.03: < 2 | [5 empty] | 0.655
| 1 | 5 | 2.02: 0.63; 2.03: < 2 | [5 empty] | 11
| 1 | 6 | 2.02: 0.95; 2.03: < 2
| Mean |  | 2.02: 0.655
| Median |  | 2.02: 0.7
| Std Dev. |  | 2.02: 0.2641
| Rel.Std.Dev. |  | 2.02: 40.32%
| PDM3 |  | 2.02: -2.22e-14%
| Z-Score (Absolute) |  | 2.02: 0.67; 2.03: 0.67
| Analytical results for Ho in OREAS 611 (Certified Value 0.22 ppm) | [24 empty] | Certified
Ho | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | [5 empty] | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | [5 empty] | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | [5 empty] | 2
2
| 1 | 1 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.24; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.24; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 8
| 1 | 3 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24 | [5 empty] | 0.21867
| 1 | 5 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.26; 2.05: 0.2; 2.06: 0.26 | [5 empty] | 32
| 1 | 6 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.26; 2.05: 0.2; 2.06: 0.24
| Mean |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24333
| Median |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.25; 2.05: 0.2; 2.06: 0.24
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 3.0405e-17; 2.04: 0.0089443; 2.05: 3.0405e-17; 2.06: 0.008165
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 1.52e-14%; 2.04: 3.578%; 2.05: 1.52e-14%; 2.06: 3.355%
| PDM3 |  | 2.02: -8.537%; 2.03: -8.537%; 2.04: 14.33%; 2.05: -8.537%; 2.06: 11.28%
Z-Score (Absolute) | bar: 2.02 to 2.06
| Analytical results for In in OREAS 611 (Certified Value 4.68 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 4.13 | 4.99 | 4.3 | 4.48 | 4.56 | 4.66 | 4.63 | 5.03 | 4.51 | 4.7 | 4.7785 | 4.76 | 4.57 | 4.5 | 5.4 | 5.05 | 4.31 | 4.75 | [5 empty] | 1
| 1 | 2 | 4.06 | 5.02 | 4.2 | 4.41 | 4.53 | 4.63 | 4.44 | 5.04 | 4.65 | 4.78 | 5.0994 | 4.73 | 4.59 | 4.66 | 5.6 | 5.07 | 4.22 | 4.76 | [5 empty] | 31
| 1 | 3 | 4.27 | 5.22 | 4.2 | 4.38 | 4.48 | 4.64 | 4.56 | 5.03 | 4.89 | 4.82 | 4.9945 | 5.11 | 4.52 | 4.68 | 5.8 | 4.89 | 4.3 | 4.75 | [5 empty] | 16
| 1 | 4 | 4.23 | 4.99 | 4.5 | 4.43 | 4.54 | 4.56 | 4.78 | 4.98 | 4.75 | 4.75 | 4.8822 | 5.15 | 4.52 | 4.7 | 5.7 | 5.03 | 4.29 | 4.7 | [5 empty] | 4.6787
| 1 | 5 | 4.25 | 4.91 | 4.6 | 4.94 | 4.57 | 4.67 | 4.74 | 5.2 | 4.47 | 4.77 | 5.0129 | 4.81 | 4.63 | 4.68 | 6 | 4.88 | 4.32 | 4.63 | [5 empty] | 33
| 1 | 6 | 4.26 | 4.94 | 4.5 | 4.81 | 4.62 | 4.66 | 4.72 | 5.19 | 4.43 | 4.92 | 4.8859 | 5.03 | 4.55 | 4.62 | 5.9 | 5.08 | 4.3 | 4.6
| Mean |  | 4.2 | 5.0117 | 4.3833 | 4.575 | 4.55 | 4.6367 | 4.645 | 5.0783 | 4.6167 | 4.79 | 4.9422 | 4.9317 | 4.5633 | 4.64 | 5.7333 | 5 | 4.29 | 4.6983
| Median |  | 4.24 | 4.99 | 4.4 | 4.455 | 4.55 | 4.65 | 4.675 | 5.035 | 4.58 | 4.775 | 4.9402 | 4.92 | 4.56 | 4.67 | 5.75 | 5.04 | 4.3 | 4.725
| Std Dev. |  | 0.085323 | 0.10944 | 0.17224 | 0.23822 | 0.046476 | 0.040332 | 0.12834 | 0.092826 | 0.17963 | 0.074833 | 0.11488 | 0.18659 | 0.04274 | 0.073756 | 0.21602 | 0.090774 | 0.035777 | 0.068532
| Rel.Std.Dev. |  | 2.031% | 2.184% | 3.929% | 5.207% | 1.021% | 0.8698% | 2.763% | 1.828% | 3.891% | 1.562% | 2.324% | 3.784% | 0.9366% | 1.59% | 3.768% | 1.815% | 0.834% | 1.459%
| PDM3 |  | -10.23% | 7.116% | -6.314% | -2.217% | -2.752% | -0.8992% | -0.7211% | 8.541% | -1.327% | 2.378% | 5.632% | 5.406% | -2.467% | -0.828% | 22.54% | 6.866% | -8.309% | 0.4188%
| Z-Score (Absolute) |  | 1.47 | 1.22 | 0.86 | 0.22 | 0.31 | 0.02 | 0.01 | 1.45 | 0.09 | 0.49 | 0.99 | 0.96 | 0.26 | 0.01 | 3.62 | 1.19 | 1.17 | 0.19
| Analytical results for K in OREAS 611 (Certified Value 1.86 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 1.8677 | 2.16 | 1.73 | 1.79 | 1.83 | 1.86 | 1.78 | 1.87 | 1.74 | 1.8868 | 1.9132 | 1.8497 | 1.91 | 1.86 | 1.9106 | 1.8 | 1.735 | 1.83 | 1.88 | 1.85 | 2 | 1.95 | 1.91 | 1
| 1 | 2 | 1.822 | 2.17 | 1.8 | 1.8 | 1.81 | 1.82 | 1.77 | 1.88 | 1.87 | 1.9383 | 1.9183 | 1.8672 | 1.93 | 1.84 | 1.8771 | 1.84 | 1.747 | 1.84 | 1.91 | 1.8 | 1.99 | 1.96 | 1.967 | #N/A
| 1 | 3 | 1.8476 | 2.28 | 1.79 | 1.84 | 1.81 | 1.82 | 1.8 | 1.85 | 1.72 | 1.9567 | 1.9188 | 1.8796 | 1.91 | 1.85 | 1.8675 | 1.83 | 1.743 | 1.8 | 1.93 | 1.84 | 2 | 1.94 | 1.984 | 16
| 1 | 4 | 1.8157 | 2.16 | 1.81 | 1.83 | 1.83 | 1.81 | 1.81 | 1.87 | 1.85 | 1.8645 | 1.8762 | 1.8645 | 1.98 | 1.86 | 1.8567 | 1.8 | 1.728 | 1.78 | 1.93 | 1.83 | 1.99 | 1.98 | 2.021 | 1.8594
| 1 | 5 | 1.8473 | 2.29 | 1.86 | 1.87 | 1.81 | 1.79 | 1.79 | 1.91 | 1.85 | 1.922 | 1.8817 | 1.8628 | 1.92 | 1.85 | 1.9018 | 1.87 | 1.727 | 1.77 | 1.93 | 1.75 | 2 | 1.95 | 1.958 | 34
| 1 | 6 | 1.8467 | 2.04 | 1.69 | 1.82 | 1.82 | 1.82 | 1.76 | 1.91 | 1.85 | 1.8949 | 1.9041 | 1.9201 | 1.9 | 1.85 | 1.9156 | 1.82 | 1.703 | 1.81 | 1.89 | 1.86 | 2 | 1.94 | 1.895
| Mean |  | 1.8412 | 2.1833 | 1.78 | 1.825 | 1.8183 | 1.82 | 1.785 | 1.8817 | 1.8133 | 1.9105 | 1.9021 | 1.874 | 1.925 | 1.8517 | 1.8882 | 1.8267 | 1.7305 | 1.805 | 1.9117 | 1.8217 | 1.9967 | 1.9533 | 1.9558
| Median |  | 1.847 | 2.165 | 1.795 | 1.825 | 1.815 | 1.82 | 1.785 | 1.875 | 1.85 | 1.9084 | 1.9087 | 1.8659 | 1.915 | 1.85 | 1.8894 | 1.825 | 1.7315 | 1.805 | 1.92 | 1.835 | 2 | 1.95 | 1.9625
| Std Dev. |  | 0.019129 | 0.092232 | 0.060663 | 0.02881 | 0.0098319 | 0.022804 | 0.018708 | 0.024014 | 0.06532 | 0.034527 | 0.018734 | 0.024532 | 0.02881 | 0.0075277 | 0.024387 | 0.026583 | 0.015643 | 0.027386 | 0.022286 | 0.040702 | 0.005164 | 0.015055 | 0.046842
| Rel.Std.Dev. |  | 1.039% | 4.224% | 3.408% | 1.579% | 0.5407% | 1.253% | 1.048% | 1.276% | 3.602% | 1.807% | 0.9849% | 1.309% | 1.497% | 0.4065% | 1.292% | 1.455% | 0.904% | 1.517% | 1.166% | 2.234% | 0.2586% | 0.7708% | 2.395%
| PDM3 |  | -0.9794% | 17.42% | -4.269% | -1.849% | -2.207% | -2.118% | -4% | 1.199% | -2.476% | 2.751% | 2.296% | 0.7856% | 3.529% | -0.4147% | 1.551% | -1.759% | -6.931% | -2.924% | 2.812% | -2.028% | 7.384% | 5.053% | 5.188%
| Z-Score (Absolute) |  | 0.15 | 4.79 | 1.04 | 0.39 | 0.48 | 0.46 | 0.96 | 0.43 | 0.55 | 0.85 | 0.73 | 0.32 | 1.06 | 0 | 0.53 | 0.36 | 1.75 | 0.67 | 0.87 | 0.43 | 2.1 | 1.47 | 1.51
| Analytical results for La in OREAS 611 (Certified Value 20.8 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 1
1
| 1 | 1 | 14.5 | 10.9 | 19.5 | 19.6 | 21.9 | 21.5 | 22.7 | 22.5 | 15.5 | 20.6 | 27.5 | 26.1907 | 18 | 17.851 | 16.2 | 21.5 | 15.8 | 26 | 16 | 24.88 |  |  |  | 1
| 1 | 2 | 15.4 | 11.3 | 20.1 | 19.5 | 21.7 | 20 | 23.2 | 23.75 | 15.9 | 23.9 | 27.3 | 26.2342 | 18.3 | 18.215 | 17 | 20.4 | 17.5 | 25 | 16 | 25.01 |  |  |  | 32
| 1 | 3 | 13.8 | 8.1 | 20.4 | 18.8 | 21.8 | 21 | 22.5 | 24.24 | 19 | 22.9 | 26.8 | 26.1023 | 18.7 | 19.189 | 16.9 | 20.9 | 18.7 | 24 | 16 | 25.14 |  |  |  | 16
| 1 | 4 | 15.9 | 7.3 | 19.7 | 19.4 | 21.9 | 23.5 | 22.4 | 24.03 | 17.2 | 24.9 | 26.9 | 26.0866 | 16.3 | 19.974 | 17.5 | 19.4 | 17.6 | 25 | 16 | 25.27 |  |  |  | 20.7995
| 1 | 5 | 15.7 | 8.7 | 22.1 | 19.2 | 21.5 | 22.2 | 24.3 | 23.82 | 15.3 | 25 | 26.6 | 25.7483 | 18.5 | 17.533 | 16.7 | 19.4 | 18.3 | 23 | 16 | 25.4 |  |  |  | 35
| 1 | 6 | 15.3 | 8.9 | 21.7 | 20.9 | 21.5 | 21.2 | 22.9 | 23.81 | 17.3 | 23.3 | 26 | 26.2611 | 18.2 | 17.671 | 17.4 | 19.9 | 17.2 | 24 | 16 | 25.53
| Mean |  | 15.1 | 9.2 | 20.5833 | 19.5667 | 21.7167 | 21.5667 | 23 | 23.6917 | 16.7 | 23.4333 | 26.85 | 26.1039 | 18 | 18.4055 | 16.95 | 20.25 | 17.5167 | 24.5 | 16 | 25.205
| Median |  | 15.35 | 8.8 | 20.25 | 19.45 | 21.75 | 21.35 | 22.8 | 23.815 | 16.55 | 23.6 | 26.85 | 26.1465 | 18.25 | 18.033 | 16.95 | 20.15 | 17.55 | 24.5 | 16 | 25.205
| Std Dev. |  | 0.7975 | 1.5786 | 1.0741 | 0.71181 | 0.18348 | 1.1877 | 0.69857 | 0.61141 | 1.41 | 1.6219 | 0.53198 | 0.1876 | 0.86718 | 0.97145 | 0.47645 | 0.84558 | 1.0068 | 1.0488 | 0 | 0.24321
| Rel.Std.Dev. |  | 5.281% | 17.16% | 5.218% | 3.638% | 0.8449% | 5.507% | 3.037% | 2.581% | 8.443% | 6.921% | 1.981% | 0.7187% | 4.818% | 5.278% | 2.811% | 4.176% | 5.748% | 4.281% | 0% | 0.9649%
| PDM3 |  | -27.4% | -55.77% | -1.039% | -5.927% | 4.409% | 3.688% | 10.58% | 13.9% | -19.71% | 12.66% | 29.09% | 25.5% | -13.46% | -11.51% | -18.51% | -2.642% | -15.78% | 17.79% | -23.08% | 21.18%
| Z-Score (Absolute) |  | 1.14 | 2.4 | 0.04 | 0.18 | 0.28 | 0.25 | 0.55 | 0.7 | 0.8 | 0.65 | 1.38 | 1.22 | 0.52 | 0.43 | 0.74 | 0.04 | 0.62 | 0.88 | 0.95 | 1.03
| Analytical results for Li in OREAS 611 (Certified Value 29.7 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 28.7 | 31.8 | 27.8 | 27.2 | 29.2 | 31.2 | 30.6 | 31.5 | 30.9 | 27.5 | 28.9969 | 30 | 29.7 | 29.5 | 32.6175 | 28.2 | 25.782 | 32 | 29 | 30 | 37.2 | 29 | 31.21 | 1
| 1 | 2 | 28.4 | 33.6 | 28.3 | 27.1 | 28.9 | 31 | 29.7 | 31.7 | 28.5 | 27.8 | 29.0602 | 30.8 | 29.9 | 30.5 | 32.1101 | 29 | 27.37 | 29 | 29 | 30 | 37.7 | 29 | 30.87 | #N/A
| 1 | 3 | 29.9 | 31.3 | 28.1 | 26.7 | 28.6 | 31 | 30 | 30.7 | 30.2 | 28.4 | 29.3829 | 31.4 | 29.9 | 30 | 32.426 | 28.9 | 26.099 | 30 | 29 | 31 | 33 | 29 | 30.53 | 16
| 1 | 4 | 28.8 | 36.4 | 28.1 | 26.8 | 29.2 | 31 | 31.8 | 31.2 | 28.4 | 27.3 | 27.9446 | 30.8 | 30.7 | 30 | 32.6195 | 28.9 | 26.421 | 31 | 29 | 30 | 35.3 | 29 | 30.19 | 29.6931
| 1 | 5 | 29.7 | 32.8 | 29.3 | 29.7 | 28.7 | 30.9 | 30.2 | 32.1 | 29.7 | 27.3 | 29.3006 | 31.5 | 29.7 | 30 | 33.6914 | 29.2 | 27.674 | 31 | 29 | 29 | 33 | 29 | 29.85 | 36
| 1 | 6 | 30.2 | 30.7 | 29.6 | 28.7 | 30.8 | 30.9 | 30.2 | 31.4 | 30.3 | 27.5 | 28.2243 | 32 | 29.3 | 29.5 | 33.2828 | 29.3 | 28.361 | 31 | 29 | 30 | 31.2 | 29 | 29.99
| Mean |  | 29.2833 | 32.7667 | 28.5333 | 27.7 | 29.2333 | 31 | 30.4167 | 31.4333 | 29.6667 | 27.6333 | 28.8183 | 31.0833 | 29.8667 | 29.9167 | 32.7912 | 28.9167 | 26.9512 | 30.6667 | 29 | 30 | 34.5667 | 29 | 30.44
| Median |  | 29.25 | 32.3 | 28.2 | 27.15 | 29.05 | 31 | 30.2 | 31.45 | 29.95 | 27.5 | 29.0286 | 31.1 | 29.8 | 30 | 32.6185 | 28.95 | 26.8955 | 31 | 29 | 30 | 34.15 | 29 | 30.36
| Std Dev. |  | 0.7414 | 2.0617 | 0.73394 | 1.2182 | 0.80664 | 0.10954 | 0.73869 | 0.47188 | 1.0172 | 0.41793 | 0.59294 | 0.69976 | 0.46332 | 0.37639 | 0.58461 | 0.38687 | 1.006 | 1.0328 | 0 | 0.63246 | 2.5897 | 0 | 0.52896
| Rel.Std.Dev. |  | 2.532% | 6.292% | 2.572% | 4.398% | 2.759% | 0.3534% | 2.429% | 1.501% | 3.429% | 1.512% | 2.058% | 2.251% | 1.551% | 1.258% | 1.783% | 1.338% | 3.733% | 3.368% | 0% | 2.108% | 7.492% | 0% | 1.738%
| PDM3 |  | -1.38% | 10.35% | -3.906% | -6.712% | -1.548% | 4.402% | 2.437% | 5.861% | -0.08889% | -6.937% | -2.946% | 4.682% | 0.5847% | 0.7531% | 10.43% | -2.615% | -9.234% | 3.279% | -2.334% | 1.034% | 16.41% | -2.334% | 2.516%
| Z-Score (Absolute) |  | 0.41 | 2.06 | 0.95 | 1.54 | 0.45 | 0.8 | 0.39 | 1.11 | 0.14 | 1.59 | 0.74 | 0.86 | 0 | 0.04 | 2.08 | 0.67 | 2.07 | 0.57 | 0.62 | 0.09 | 3.34 | 0.62 | 0.41
| Analytical results for Lu in OREAS 611 (Indicative Value 86.7 ppb) | [24 empty] | Indicative
Lu | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16; 2.07: 2.20; 2.08: 2.21 | ppb
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*OES/MS | 1
1
| 1 | 1 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100; 2.07: 80; 2.08: 90 | 1
| 1 | 2 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100; 2.07: 80; 2.08: 90 | 6
| 1 | 3 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100; 2.07: 70; 2.08: 90 | 16
| 1 | 4 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100; 2.07: 70; 2.08: 80 | 86.6667
| 1 | 5 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100; 2.07: 80; 2.08: 80 | 12
| 1 | 6 | 2.02: < 100; 2.03: < 100; 2.04: 80; 2.05: < 100; 2.06: 100; 2.07: 80; 2.08: 80
| Mean |  | 2.04: 85; 2.06: 100; 2.07: 76.6667; 2.08: 85
| Median |  | 2.04: 85; 2.06: 100; 2.07: 80; 2.08: 85
| Std Dev. |  | 2.04: 5.4772; 2.06: 0; 2.07: 5.164; 2.08: 5.4772
| Rel.Std.Dev. |  | 2.04: 6.444%; 2.06: 0%; 2.07: 6.736%; 2.08: 6.444%
| PDM3 |  | 2.04: -1.923%; 2.06: 15.38%; 2.07: -11.54%; 2.08: -1.923%
| Z-Score (Absolute) |  | 2.02: 0.77; 2.03: 0.77; 2.04: 0.24; 2.05: 0.77; 2.06: 0.67; 2.07: 0; 2.08: 0.24
| Analytical results for Mg in OREAS 611 (Certified Value 1472 ppm) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 1517 | 1400 | 1400 | 1300 | 1400 | 1400 | 1400 | 1400 | 1500 | 1490 | 1474.1116 | 1372 | 1600 | 1500 | 1617.1544 | 1500 | 1280 | 1500 | 1600 | 1500 | 1600 | 1100 | 1554 | 1
| 1 | 2 | 1478 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1590 | 1539 | 1517.4026 | 1470 | 1600 | 1500 | 1591.4389 | 1500 | 1290 | 1400 | 1500 | 1400 | 1600 | 1200 | 1582 | #N/A
| 1 | 3 | 1497 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1610 | 1551 | 1533.2545 | 1506 | 1600 | 1600 | 1582.7288 | 1500 | 1280 | 1300 | 1500 | 1500 | 1600 | 1200 | 1610 | 16
| 1 | 4 | 1464 | 1500 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1560 | 1478 | 1482.7921 | 1434 | 1600 | 1600 | 1576.0002 | 1500 | 1290 | 1400 | 1500 | 1400 | 1500 | 1200 | 1638 | 1472.4424
| 1 | 5 | 1496 | 1400 | 1500 | 1400 | 1400 | 1300 | 1400 | 1500 | 1580 | 1521 | 1510.9867 | 1461 | 1600 | 1500 | 1610.2974 | 1500 | 1210 | 1300 | 1500 | 1400 | 1500 | 1100 | 1666 | 37
| 1 | 6 | 1519 | 1300 | 1500 | 1400 | 1400 | 1300 | 1400 | 1500 | 1540 | 1497 | 1505.5376 | 1518 | 1600 | 1500 | 1624.6873 | 1500 | 1280 | 1400 | 1500 | 1500 | 1500 | 1200 | 1694
| Mean |  | 1495.1667 | 1400 | 1433.3333 | 1383.3333 | 1400 | 1366.6667 | 1400 | 1433.3333 | 1563.3333 | 1512.6667 | 1504.0142 | 1460.1667 | 1600 | 1533.3333 | 1600.3845 | 1500 | 1271.6667 | 1383.3333 | 1516.6667 | 1450 | 1550 | 1166.6667 | 1624
| Median |  | 1496.5 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1400 | 1570 | 1509 | 1508.2621 | 1465.5 | 1600 | 1500 | 1600.8682 | 1500 | 1280 | 1400 | 1500 | 1450 | 1550 | 1200 | 1624
| Std Dev. |  | 21.5167 | 63.2456 | 51.6398 | 40.8248 | 0 | 51.6398 | 0 | 51.6398 | 39.3277 | 28.9597 | 22.0442 | 52.9166 | 0 | 51.6398 | 19.7811 | 0 | 30.605 | 75.2773 | 40.8248 | 54.7723 | 54.7723 | 51.6398 | 52.3832
| Rel.Std.Dev. |  | 1.439% | 4.518% | 3.603% | 2.951% | 0% | 3.779% | 0% | 3.603% | 2.516% | 1.914% | 1.466% | 3.624% | 0% | 3.368% | 1.236% | 0% | 2.407% | 5.442% | 2.692% | 3.777% | 3.534% | 4.426% | 3.226%
| PDM3 |  | 1.543% | -4.92% | -2.656% | -6.052% | -4.92% | -7.184% | -4.92% | -2.656% | 6.173% | 2.732% | 2.144% | -0.8337% | 8.663% | 4.135% | 8.689% | 1.872% | -13.64% | -6.052% | 3.003% | -1.524% | 5.267% | -20.77% | 10.29%
| Z-Score (Absolute) |  | 0.39 | 0.67 | 0.3 | 0.86 | 0.67 | 1.05 | 0.67 | 0.3 | 1.16 | 0.59 | 0.49 | 0 | 1.57 | 0.82 | 1.57 | 0.45 | 2.11 | 0.86 | 0.63 | 0.11 | 1.01 | 3.29 | 1.84
| Analytical results for Mn in OREAS 611 (Certified Value 79 ppm) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 80 | 85 | 75 | 76 | 78 | 80 | 78 | 80 | 65 | 82 | 78.8672 | 75 | 83 | 84 | 79.2669 | 73 | 75.298 | 75 | 85 | 75 | 81 | 96 | 77.2 | 1
| 1 | 2 | 77 | 86 | 77 | 77 | 78 | 77 | 79 | 80 | 68 | 84 | 77.8354 | 78 | 82 | 82 | 78.8748 | 76 | 76.122 | 74 | 85 | 73 | 82 | 97 | 76.61 | 34
| 1 | 3 | 78 | 88 | 80 | 80 | 77 | 77 | 78 | 79 | 67 | 84 | 81.3755 | 77 | 79 | 84 | 79.1802 | 79 | 74.689 | 71 | 83 | 77 | 79 | 97 | 76.02 | 16
| 1 | 4 | 75 | 84 | 76 | 77 | 79 | 77 | 80 | 79 | 68 | 82 | 75.7153 | 75 | 80 | 88 | 77.5498 | 75 | 76.33 | 70 | 83 | 75 | 81 | 97 | 76.43 | 78.8582
| 1 | 5 | 78 | 87 | 70 | 80 | 79 | 78 | 77 | 83 | 69 | 83 | 79.7346 | 75 | 81 | 84 | 78.4959 | 78 | 72.499 | 71 | 85 | 72 | 80 | 97 | 76.84 | 38
| 1 | 6 | 78 | 78 | 82 | 77 | 78 | 79 | 78 | 82 | 69 | 82 | 77.2863 | 81 | 80 | 88 | 79.1515 | 78 | 75.437 | 70 | 85 | 79 | 79 | 96 | 75.25
| Mean |  | 77.6667 | 84.6667 | 76.6667 | 77.8333 | 78.1667 | 78 | 78.3333 | 80.5 | 67.6667 | 82.8333 | 78.4691 | 76.8333 | 80.8333 | 85 | 78.7532 | 76.5 | 75.0625 | 71.8333 | 84.3333 | 75.1667 | 80.3333 | 96.6667 | 76.3917
| Median |  | 78 | 85.5 | 76.5 | 77 | 78 | 77.5 | 78 | 80 | 68 | 82.5 | 78.3513 | 76 | 80.5 | 84 | 79.0132 | 77 | 75.3675 | 71 | 85 | 75 | 80.5 | 97 | 76.52
| Std Dev. |  | 1.633 | 3.559 | 4.1793 | 1.7224 | 0.75277 | 1.2649 | 1.0328 | 1.6432 | 1.5055 | 0.98319 | 1.9795 | 2.4014 | 1.472 | 2.4495 | 0.65333 | 2.2583 | 1.3882 | 2.137 | 1.0328 | 2.5626 | 1.2111 | 0.5164 | 0.68485
| Rel.Std.Dev. |  | 2.103% | 4.204% | 5.451% | 2.213% | 0.963% | 1.622% | 1.318% | 2.041% | 2.225% | 1.187% | 2.523% | 3.125% | 1.821% | 2.882% | 0.8296% | 2.952% | 1.849% | 2.975% | 1.225% | 3.409% | 1.508% | 0.5342% | 0.8965%
| PDM3 |  | -1.511% | 7.366% | -2.779% | -1.3% | -0.8769% | -1.088% | -0.6656% | 2.082% | -14.19% | 5.041% | -0.4935% | -2.568% | 2.505% | 7.788% | -0.1332% | -2.99% | -4.813% | -8.908% | 6.943% | -4.681% | 1.871% | 22.58% | -3.128%
| Z-Score (Absolute) |  | 0.16 | 2.02 | 0.47 | 0.1 | 0 | 0.05 | 0.05 | 0.73 | 3.27 | 1.45 | 0.09 | 0.41 | 0.83 | 2.13 | 0.18 | 0.52 | 0.97 | 1.97 | 1.92 | 0.93 | 0.67 | 5.76 | 0.55
| Analytical results for Mo in OREAS 611 (Certified Value 5.46 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 5.1 | 5.04 | 4.98 | 5.47 | 6.8 | 5.58 | 6 | 5.65 | 5.8 | 5.5 | 5.3544 | 5.82 | 6.37 | 5.5 | 8.1835 | 4.72 | 5.042 | 5.87 | 5.23 | 5 | 5.6 | 4 | 8.713 | 1
| 1 | 2 | 4.9 | 5.67 | 4.97 | 5.33 | 5.36 | 5.74 | 5.53 | 5.39 | 5.7 | 5.8 | 5.6083 | 4.8 | 6.33 | 5.5 | 7.925 | 5.34 | 5.405 | 5.86 | 5.11 | 5 | 5.8 | 4 | 8.648 | 35
| 1 | 3 | 5 | 5.29 | 5.1 | 5.58 | 5.85 | 5.46 | 5.46 | 5.79 | 5.6 | 5.9 | 5.7354 | 4.86 | 6.05 | 5.5 | 8.4372 | 5.16 | 4.512 | 5.82 | 5.07 | 5 | 5.9 | 5 | 8.683 | 16
| 1 | 4 | 5.2 | 5.42 | 5.32 | 5.22 | 5.44 | 5.41 | 5.6 | 5.29 | 5.7 | 6 | 5.2607 | 4.95 | 6.4 | 5.5 | 7.9077 | 5.13 | 4.981 | 5.69 | 4.98 | 6 | 5.9 | 4 | 8.618 | 5.4576
| 1 | 5 | 5 | 5.56 | 5.21 | 5.78 | 5.35 | 5.16 | 6.11 | 5.65 | 5.7 | 5.6 | 5.5409 | 5.16 | 6.27 | 5.5 | 7.8144 | 5.11 | 5.291 | 5.71 | 4.99 | 5 | 5.7 | 5 | 8.553 | 39
| 1 | 6 | 5 | 5.08 | 5.03 | 5.3 | 5.41 | 5.18 | 5.61 | 5.73 | 5.6 | 5.2 | 5.4965 | 5.71 | 6.12 | 5.5 | 8.5307 | 4.34 | 4.922 | 5.77 | 5.14 | 5 | 5.6 | 4 | 8.588
| Mean |  | 5.0333 | 5.3433 | 5.1017 | 5.4467 | 5.7017 | 5.4217 | 5.7183 | 5.5833 | 5.6833 | 5.6667 | 5.4994 | 5.2167 | 6.2567 | 5.5 | 8.1331 | 4.9667 | 5.0255 | 5.7867 | 5.0867 | 5.1667 | 5.75 | 4.3333 | 8.6338
| Median |  | 5 | 5.355 | 5.065 | 5.4 | 5.425 | 5.435 | 5.605 | 5.65 | 5.7 | 5.7 | 5.5187 | 5.055 | 6.3 | 5.5 | 8.0543 | 5.12 | 5.0115 | 5.795 | 5.09 | 5 | 5.75 | 4 | 8.633
| Std Dev. |  | 0.10328 | 0.25445 | 0.13906 | 0.20781 | 0.56961 | 0.22578 | 0.26858 | 0.19826 | 0.075277 | 0.29439 | 0.17165 | 0.44329 | 0.14166 | 0 | 0.29961 | 0.36811 | 0.31342 | 0.07607 | 0.094798 | 0.40825 | 0.13784 | 0.5164 | 0.059617
| Rel.Std.Dev. |  | 2.052% | 4.762% | 2.726% | 3.815% | 9.99% | 4.164% | 4.697% | 3.551% | 1.325% | 5.195% | 3.121% | 8.498% | 2.264% | 0% | 3.684% | 7.412% | 6.237% | 1.315% | 1.864% | 7.902% | 2.397% | 11.92% | 0.6905%
| PDM3 |  | -7.773% | -2.093% | -6.521% | -0.1998% | 4.473% | -0.6579% | 4.778% | 2.304% | 4.137% | 3.831% | 0.7656% | -4.414% | 14.64% | 0.7774% | 49.02% | -8.995% | -7.917% | 6.03% | -6.796% | -5.33% | 5.358% | -20.6% | 58.2%
| Z-Score (Absolute) |  | 1.26 | 0.42 | 1.07 | 0.14 | 0.54 | 0.21 | 0.59 | 0.22 | 0.49 | 0.45 | 0 | 0.76 | 2.04 | 0 | 7.1 | 1.44 | 1.28 | 0.77 | 1.11 | NA | 0.67 | NA | 8.45
Results from laboratories 22& 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 611 (Certified Value 0.804 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.8174 | 0.824 | 0.81 | 0.77 | 0.79 | 0.8 | 0.77 | 0.79 | 0.838 | 0.7965 | 0.78487 | 0.8077 | 0.82 | 0.85 | 0.93164 | 0.82 | 0.73 | 0.84 | 0.75 | 0.79 | 1.07 | 0.81 | 0.895 | 1
| 1 | 2 | 0.7948 | 0.833 | 0.8 | 0.77 | 0.78 | 0.79 | 0.77 | 0.8 | 0.839 | 0.8166 | 0.79606 | 0.7769 | 0.83 | 0.85 | 0.87292 | 0.8 | 0.726 | 0.84 | 0.78 | 0.77 | 1.07 | 0.81 | 0.8651 | #N/A
| 1 | 3 | 0.8005 | 0.808 | 0.79 | 0.79 | 0.78 | 0.78 | 0.77 | 0.79 | 0.841 | 0.8243 | 0.81826 | 0.7881 | 0.83 | 0.86 | 0.83597 | 0.81 | 0.731 | 0.82 | 0.76 | 0.79 | 1.08 | 0.82 | 0.8952 | 16
| 1 | 4 | 0.7903 | 0.864 | 0.81 | 0.78 | 0.79 | 0.78 | 0.78 | 0.8 | 0.835 | 0.7878 | 0.77034 | 0.7821 | 0.85 | 0.86 | 0.83007 | 0.8 | 0.725 | 0.81 | 0.75 | 0.78 | 1.06 | 0.83 | 0.89453 | 0.80378
| 1 | 5 | 0.8138 | 0.831 | 0.82 | 0.81 | 0.77 | 0.74 | 0.77 | 0.83 | 0.83 | 0.8142 | 0.80091 | 0.7795 | 0.82 | 0.85 | 0.91816 | 0.81 | 0.71 | 0.81 | 0.79 | 0.75 | 1.07 | 0.82 | 0.89954 | 40
| 1 | 6 | 0.8094 | 0.763 | 0.88 | 0.78 | 0.77 | 0.75 | 0.76 | 0.81 | 0.833 | 0.8032 | 0.76989 | 0.8497 | 0.82 | 0.85 | 0.9211 | 0.8 | 0.706 | 0.81 | 0.75 | 0.79 | 1.04 | 0.81 | 0.91
| Mean |  | 0.80437 | 0.8205 | 0.81833 | 0.78333 | 0.78 | 0.77333 | 0.77 | 0.80333 | 0.836 | 0.8071 | 0.79006 | 0.79733 | 0.82833 | 0.85333 | 0.88498 | 0.80667 | 0.72133 | 0.82167 | 0.76333 | 0.77833 | 1.065 | 0.81667 | 0.89323
| Median |  | 0.80495 | 0.8275 | 0.81 | 0.78 | 0.78 | 0.78 | 0.77 | 0.8 | 0.8365 | 0.8087 | 0.79046 | 0.7851 | 0.825 | 0.85 | 0.89554 | 0.805 | 0.7255 | 0.815 | 0.755 | 0.785 | 1.07 | 0.815 | 0.8951
| Std Dev. |  | 0.010849 | 0.033566 | 0.031885 | 0.015055 | 0.0089443 | 0.023381 | 0.0063246 | 0.015055 | 0.0040988 | 0.013686 | 0.018822 | 0.027938 | 0.01169 | 0.005164 | 0.045046 | 0.008165 | 0.010652 | 0.01472 | 0.017512 | 0.016021 | 0.013784 | 0.008165 | 0.014974
| Rel.Std.Dev. |  | 1.349% | 4.091% | 3.896% | 1.922% | 1.147% | 3.023% | 0.8214% | 1.874% | 0.4903% | 1.696% | 2.382% | 3.504% | 1.411% | 0.6052% | 5.09% | 1.012% | 1.477% | 1.791% | 2.294% | 2.058% | 1.294% | 0.9998% | 1.676%
| PDM3 |  | 0.07339% | 2.081% | 1.811% | -2.543% | -2.958% | -3.788% | -4.202% | -0.05517% | 4.009% | 0.4135% | -1.707% | -0.8016% | 3.055% | 6.165% | 10.1% | 0.3595% | -10.26% | 2.226% | -5.032% | -3.165% | 32.5% | 1.604% | 11.13%
| Z-Score (Absolute) |  | 0.07 | 0.4 | 0.34 | 0.67 | 0.77 | 0.96 | 1.06 | 0.1 | 0.85 | 0.01 | 0.48 | 0.27 | 0.63 | 1.35 | 2.26 | 0 | 2.47 | 0.43 | 1.25 | 0.82 | 7.47 | 0.29 | 2.5
| Analytical results for Nb in OREAS 611 (Certified Value 8.82 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 8.2 | 8.47 | 7.2 | 9 | 8.7 | 8.8 | 9.5 | 10.3 | 8.3 | 8.69 | 9.4181 | 9.5 | 11.1 | 9 | 7.6751 | 8.6 | 8.4 | 9 | 9.6 | 7 |  |  |  | 1
| 1 | 2 | 8.2 | 8.82 | 7.5 | 9.2 | 8.7 | 8.6 | 9.1 | 10.2 | 8 | 8.95 | 9.3076 | 9.04 | 10.6 | 9 | 7.6976 | 9 | 8.3 | 8.7 | 9.5 | 7 |  |  |  | 36
| 1 | 3 | 8.3 | 8.88 | 7.7 | 9.3 | 8.6 | 8.6 | 9.3 | 9.7 | 8.3 | 8.99 | 8.7556 | 9.42 | 10.7 | 9 | 7.646 | 8.9 | 8.6 | 8.6 | 9.7 | 7 |  |  |  | 16
| 1 | 4 | 8.2 | 9.17 | 7.4 | 9.3 | 8.6 | 8.5 | 9.4 | 10 | 7.9 | 8.77 | 8.4794 | 8.57 | 10.6 | 9 | 7.6237 | 8.7 | 8.2 | 8.6 | 9.5 | 7 |  |  |  | 8.8228
| 1 | 5 | 8.5 | 9.1 | 6.2 | 9.5 | 8.9 | 8.5 | 9.4 | 10.1 | 8.4 | 8.73 | 8.4162 | 9.07 | 10.5 | 8.5 | 7.6478 | 9.1 | 8.2 | 8.7 | 9.3 | 7 |  |  |  | 41
| 1 | 6 | 8.4 | 8.32 | 6.7 | 9 | 8.8 | 8.6 | 9.1 | 10.1 | 8.1 | 8.88 | 8.7975 | 9.06 | 10.4 | 9 | 7.73 | 8.6 | 8.3 | 8.8 | 9.2 | 6
| Mean |  | 8.3 | 8.7933 | 7.1167 | 9.2167 | 8.7167 | 8.6 | 9.3 | 10.0667 | 8.1667 | 8.835 | 8.8624 | 9.11 | 10.65 | 8.9167 | 7.67 | 8.8167 | 8.3333 | 8.7333 | 9.4667 | 6.8333
| Median |  | 8.25 | 8.85 | 7.3 | 9.25 | 8.7 | 8.6 | 9.35 | 10.1 | 8.2 | 8.825 | 8.7766 | 9.065 | 10.6 | 9 | 7.6614 | 8.8 | 8.3 | 8.7 | 9.5 | 7
| Std Dev. |  | 0.12649 | 0.33845 | 0.56362 | 0.19408 | 0.1169 | 0.10954 | 0.16733 | 0.20656 | 0.19664 | 0.12292 | 0.41676 | 0.3313 | 0.2429 | 0.20412 | 0.038983 | 0.2137 | 0.15055 | 0.15055 | 0.18619 | 0.40825
| Rel.Std.Dev. |  | 1.524% | 3.849% | 7.92% | 2.106% | 1.341% | 1.274% | 1.799% | 2.052% | 2.408% | 1.391% | 4.703% | 3.637% | 2.281% | 2.289% | 0.5082% | 2.424% | 1.807% | 1.724% | 1.967% | 5.974%
| PDM3 |  | -5.925% | -0.3339% | -19.34% | 4.464% | -1.203% | -2.525% | 5.409% | 14.1% | -7.437% | 0.1384% | 0.4489% | 3.255% | 20.71% | 1.064% | -13.07% | -0.0694% | -5.548% | -1.014% | 7.298% | -22.55%
| Z-Score (Absolute) |  | 0.87 | 0.04 | 2.87 | 0.67 | 0.17 | 0.37 | 0.81 | 2.11 | 1.1 | 0.03 | 0.08 | 0.49 | 3.09 | 0.17 | 1.93 | 0 | 0.81 | 0.14 | 1.1 | NA
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 611 (Certified Value 19.4 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 |  |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS |  |  |  | 1
2
| 1 | 1 | 2.02: 18.5; 2.03: 17.4; 2.04: 19.67; 2.05: 17; 2.06: 20.9 |  |  |  | 1
| 1 | 2 | 2.02: 20.7; 2.03: 18.4; 2.04: 20.45; 2.05: 19; 2.06: 21 |  |  |  | 10
| 1 | 3 | 2.02: 19.1; 2.03: 18.4; 2.04: 20.56; 2.05: 19.8; 2.06: 20.5 |  |  |  | 16
| 1 | 4 | 2.02: 19.5; 2.03: 18.3; 2.04: 20.21; 2.05: 18.5; 2.06: 20.9 |  |  |  | 19.393
| 1 | 5 | 2.02: 19.2; 2.03: 18.8; 2.04: 20.97; 2.05: 17.6; 2.06: 20.2 |  |  |  | 42
| 1 | 6 | 2.02: 17.5; 2.03: 19.9; 2.04: 20.63; 2.05: 17.7; 2.06: 20.5
| Mean |  | 2.02: 19.0833; 2.03: 18.5333; 2.04: 20.415; 2.05: 18.2667; 2.06: 20.6667
| Median |  | 2.02: 19.15; 2.03: 18.4; 2.04: 20.505; 2.05: 18.1; 2.06: 20.7
| Std Dev. |  | 2.02: 1.0629; 2.03: 0.81404; 2.04: 0.44126; 2.05: 1.0309; 2.06: 0.31411
| Rel.Std.Dev. |  | 2.02: 5.57%; 2.03: 4.392%; 2.04: 2.161%; 2.05: 5.643%; 2.06: 1.52%
| PDM3 |  | 2.02: -1.597%; 2.03: -4.433%; 2.04: 5.27%; 2.05: -5.808%; 2.06: 6.568%
| Z-Score (Absolute) |  | 2.02: 0; 2.03: 0.45; 2.04: 1.1; 2.05: 0.67; 2.06: 1.31
| Analytical results for Ni in OREAS 611 (Certified Value 27 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 28 | 27.8 | 26.3 | 25.6 | 27.1 | 26 | 27.6 | 27 | 30 | 26.9 | 26.8784 | 27.3 | 27.1 | 34 | 27.1172 | 25.9 | 26.578 | 29.4 | 27.1 | 26 | 28.5 | 28 | 33.68 | 1
| 1 | 2 | 27 | 28.6 | 26.3 | 25.4 | 27.3 | 25.8 | 26.5 | 26.7 | 32 | 27.1 | 27.2098 | 25.2 | 27.9 | 38 | 27.5758 | 25.5 | 27.313 | 29.6 | 27.6 | 25 | 28.3 | 28 | 32.48 | 21
| 1 | 3 | 27 | 28.9 | 27.6 | 25.3 | 26.9 | 25.8 | 27.4 | 26 | 31 | 27.6 | 27.0392 | 27.3 | 27.4 | 34 | 25.3019 | 25.9 | 25.692 | 28.6 | 27.5 | 26 | 28.7 | 28 | 33.28 | 16
| 1 | 4 | 28 | 27.5 | 29 | 24.8 | 27.2 | 25.5 | 27.4 | 26.6 | 31 | 27.5 | 26.2817 | 24.8 | 26.3 | 36 | 24.6733 | 25.6 | 24.71 | 28 | 27.3 | 26 | 29 | 28 | 32.08 | 27.0308
| 1 | 5 | 27 | 28 | 27.8 | 28.1 | 27.3 | 26.9 | 26.8 | 27.3 | 32 | 27.3 | 27.3037 | 26.2 | 26.9 | 36 | 27.5396 | 26.2 | 26.606 | 27.7 | 26.8 | 25 | 29.3 | 28 | 31.88 | 43
| 1 | 6 | 28 | 25.8 | 28.8 | 26.7 | 26.8 | 27.2 | 26.3 | 27.8 | 32 | 28 | 26.4666 | 28.1 | 27.2 | 32 | 26.0866 | 25.9 | 26.481 | 28.1 | 27.2 | 26 | 29.7 | 28 | 31.68
| Mean |  | 27.5 | 27.7667 | 27.6333 | 25.9833 | 27.1 | 26.2 | 27 | 26.9 | 31.3333 | 27.4 | 26.8632 | 26.4833 | 27.1333 | 35 | 26.3824 | 25.8333 | 26.23 | 28.5667 | 27.25 | 25.6667 | 28.9167 | 28 | 32.5133
| Median |  | 27.5 | 27.9 | 27.7 | 25.5 | 27.15 | 25.9 | 27.1 | 26.85 | 31.5 | 27.4 | 26.9588 | 26.75 | 27.15 | 35 | 26.6019 | 25.9 | 26.5295 | 28.35 | 27.25 | 26 | 28.85 | 28 | 32.28
| Std Dev. |  | 0.54772 | 1.093 | 1.1673 | 1.2123 | 0.20976 | 0.68411 | 0.54037 | 0.61968 | 0.8165 | 0.38987 | 0.41008 | 1.3045 | 0.53166 | 2.0976 | 1.2231 | 0.25033 | 0.90534 | 0.78145 | 0.2881 | 0.5164 | 0.52313 | 0 | 0.80416
| Rel.Std.Dev. |  | 1.992% | 3.936% | 4.224% | 4.666% | 0.774% | 2.611% | 2.001% | 2.304% | 2.606% | 1.423% | 1.527% | 4.926% | 1.959% | 5.993% | 4.636% | 0.969% | 3.452% | 2.736% | 1.057% | 2.012% | 1.809% | 0% | 2.473%
| PDM3 |  | 1.736% | 2.722% | 2.229% | -3.875% | 0.2561% | -3.073% | -0.1138% | -0.4838% | 15.92% | 1.366% | -0.6198% | -2.025% | 0.3795% | 29.48% | -2.399% | -4.43% | -2.962% | 5.682% | 0.8111% | -5.046% | 6.977% | 3.586% | 20.28%
| Z-Score (Absolute) |  | NA | 0.77 | 0.63 | 1.15 | 0.05 | 0.92 | 0.05 | 0.16 | NA | 0.38 | 0.2 | 0.61 | 0.09 | NA | 0.72 | 1.31 | 0.88 | 1.64 | 0.22 | NA | 2.01 | NA | 5.89
Results from laboratories 2, 10, 16, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 611 (Certified Value 548 ppm) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 549 | 530 | 550 | 540 | 520 | 540 | 520 | 540 | 570 | 559 | 526.0162 | 541 | 538 | 550 | 559.7919 | 700 | 690 | 698 | 600 | 800 | 590 | 690 | 544.4 | 1
| 1 | 2 | 529 | 520 | 540 | 550 | 520 | 540 | 530 | 550 | 600 | 558 | 547.1849 | 540 | 541 | 550 | 561.3508 | 700 | 630 | 702 | 600 | 800 | 590 | 690 | 529.6 | #N/A
| 1 | 3 | 550 | 550 | 550 | 530 | 520 | 520 | 530 | 540 | 590 | 584 | 563.921 | 550 | 557 | 550 | 546.5471 | 700 | 670 | 688 | 600 | 800 | 600 | 700 | 527.48 | 16
| 1 | 4 | 547 | 520 | 570 | 540 | 530 | 540 | 540 | 560 | 590 | 561 | 544.3478 | 543 | 542 | 550 | 556.278 | 700 | 630 | 684 | 600 | 800 | 610 | 690 | 540 | 547.8208
| 1 | 5 | 536 | 530 | 570 | 570 | 520 | 520 | 520 | 560 | 600 | 569 | 547.24 | 549 | 538 | 550 | 562.5965 | 600 | 650 | 681 | 600 | 800 | 600 | 680 | 520 | 44
| 1 | 6 | 550 | 490 | 560 | 530 | 520 | 540 | 520 | 570 | 570 | 568 | 524.0203 | 569 | 536 | 550 | 557.5474 | 700 | 660 | 686 | 600 | 800 | 560 | 690 | 520.4
| Mean |  | 543.5 | 523.3333 | 556.6667 | 543.3333 | 521.6667 | 533.3333 | 526.6667 | 553.3333 | 586.6667 | 566.5 | 542.1217 | 548.6667 | 542 | 550 | 557.3519 | 683.3333 | 655 | 689.8333 | 600 | 800 | 591.6667 | 690 | 530.3133
| Median |  | 548 | 525 | 555 | 540 | 520 | 540 | 525 | 555 | 590 | 564.5 | 545.7663 | 546 | 539.5 | 550 | 558.6696 | 700 | 655 | 687 | 600 | 800 | 595 | 690 | 528.54
| Std Dev. |  | 8.8713 | 19.6638 | 12.1106 | 15.0555 | 4.0825 | 10.328 | 8.165 | 12.1106 | 13.6626 | 9.7314 | 14.9603 | 10.7827 | 7.6681 | 0 | 5.786 | 40.8248 | 23.4521 | 8.3046 | 0 | 0 | 17.224 | 6.3246 | 10.0543
| Rel.Std.Dev. |  | 1.632% | 3.757% | 2.176% | 2.771% | 0.7826% | 1.936% | 1.55% | 2.189% | 2.329% | 1.718% | 2.76% | 1.965% | 1.415% | 0% | 1.038% | 5.974% | 3.58% | 1.204% | 0% | 0% | 2.911% | 0.9166% | 1.896%
| PDM3 |  | -0.7887% | -4.47% | 1.615% | -0.8191% | -4.774% | -2.645% | -3.862% | 1.006% | 7.091% | 3.41% | -1.04% | 0.1544% | -1.063% | 0.3978% | 1.74% | 24.74% | 19.56% | 25.92% | 9.525% | 46.03% | 8.004% | 25.95% | -3.196%
| Z-Score (Absolute) |  | 0.24 | 1.06 | 0.3 | 0.24 | 1.12 | 0.65 | 0.92 | 0.16 | 1.52 | 0.7 | 0.29 | 0.03 | 0.3 | 0.03 | 0.33 | NA | 4.3 | 5.71 | NA | NA | 1.72 | 5.72 | 0.77
Results from laboratories 18, 21 & 22 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 611 (Certified Value 640 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 660 | 608.16 | 573 | 633 | 651 | 662 | 658 | 668 | 651 | 642.8 | 634.2495 | 670 | 651 | 670 | 626.1019 | 629 | 627.293 | 573.7 | 569 | 598 | 620.2 | 590 | 668.9 | 1
| 1 | 2 | 638 | 644.5 | 585 | 637 | 642 | 645 | 660 | 682 | 668 | 658.4 | 620.207 | 677.48 | 661 | 671 | 642.0866 | 614.3 | 621.706 | 578.1 | 554 | 563 | 612.4 | 580 | 681.2 | 23
| 1 | 3 | 658 | 644.96 | 604 | 650 | 645 | 643 | 671 | 672 | 659 | 654.3 | 643.0749 | 713.63 | 656 | 692 | 637.7889 | 615.1 | 639.237 | 574.5 | 561 | 560 | 585.5 | 600 | 673.5 | 16
| 1 | 4 | 652 | 657.74 | 577 | 639 | 658 | 648 | 672 | 676 | 658 | 644.6 | 612.0629 | 677.02 | 658 | 686 | 630.8637 | 612 | 626.203 | 576.6 | 577 | 578 | 574 | 590 | 675.8 | 639.5673
| 1 | 5 | 627 | 638.57 | 561 | 652 | 643 | 641 | 658 | 698 | 656 | 660.7 | 634.4883 | 718.82 | 684 | 675 | 633.1635 | 616.5 | 652.373 | 580.5 | 537 | 563 | 562.1 | 590 | 671.1 | 45
| 1 | 6 | 653 | 602.14 | 622 | 639 | 652 | 660 | 652 | 691 | 656 | 661.5 | 617.5311 | 741 | 656 | 667 | 631.6383 | 619.7 | 628.896 | 580.6 | 556 | 556 | 569.8 | 590 | 673.4
| Mean |  | 648 | 632.6783 | 587 | 641.6667 | 648.5 | 649.8333 | 661.8333 | 681.1667 | 658 | 653.7167 | 626.9356 | 699.6583 | 661 | 676.8333 | 633.6072 | 617.7667 | 632.618 | 577.3333 | 559 | 569.6667 | 587.3333 | 590 | 673.9833
| Median |  | 652.5 | 641.535 | 581 | 639 | 648 | 646.5 | 659 | 679 | 657 | 656.35 | 627.2283 | 695.555 | 657 | 673 | 632.4009 | 615.8 | 628.0945 | 577.35 | 558.5 | 563 | 579.75 | 590 | 673.45
| Std Dev. |  | 12.853 | 22.303 | 22.3159 | 7.5807 | 6.2209 | 8.9759 | 7.9603 | 11.5311 | 5.6214 | 8.172 | 12.0493 | 28.8288 | 11.7303 | 9.9482 | 5.6092 | 6.0649 | 11.2828 | 2.9344 | 13.755 | 15.7565 | 23.8069 | 6.3246 | 4.2452
| Rel.Std.Dev. |  | 1.983% | 3.525% | 3.802% | 1.181% | 0.9593% | 1.381% | 1.203% | 1.693% | 0.8543% | 1.25% | 1.922% | 4.12% | 1.775% | 1.47% | 0.8853% | 0.9817% | 1.784% | 0.5083% | 2.461% | 2.766% | 4.053% | 1.072% | 0.6299%
| PDM3 |  | 1.319% | -1.077% | -8.219% | 0.3282% | 1.397% | 1.605% | 3.481% | 6.504% | 2.882% | 2.212% | -1.975% | 9.396% | 3.351% | 5.827% | -0.9319% | -3.409% | -1.087% | -9.731% | -12.6% | -10.93% | -8.167% | -7.75% | 5.381%
| Z-Score (Absolute) |  | 0.21 | 0.3 | 1.83 | 0 | 0.23 | 0.27 | 0.67 | 1.32 | 0.55 | 0.4 | 0.49 | 1.94 | 0.65 | 1.18 | 0.27 | 0.8 | 0.3 | 2.15 | 2.76 | 2.41 | 1.82 | 1.73 | 1.08
| Analytical results for Pr in OREAS 611 (Certified Value 5.15 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | 2
2
| 1 | 1 | 2.02: 4.9; 2.03: 4.6; 2.04: 5.5; 2.05: 4.3; 2.06: 5.9 | 1
| 1 | 2 | 2.02: 5.3; 2.03: 4.6; 2.04: 5.67; 2.05: 4.7; 2.06: 6.02 | 12
| 1 | 3 | 2.02: 4.8; 2.03: 4.4; 2.04: 5.79; 2.05: 5.2; 2.06: 5.84 | 16
| 1 | 4 | 2.02: 5.2; 2.03: 4.4; 2.04: 5.65; 2.05: 4.8; 2.06: 5.96 | 5.1463
| 1 | 5 | 2.02: 5.1; 2.03: 4.5; 2.04: 5.77; 2.05: 4.7; 2.06: 5.82 | 46
| 1 | 6 | 2.02: 4.4; 2.03: 4.7; 2.04: 5.79; 2.05: 4.4; 2.06: 5.68
| Mean |  | 2.02: 4.95; 2.03: 4.5333; 2.04: 5.695; 2.05: 4.6833; 2.06: 5.87
| Median |  | 2.02: 5; 2.03: 4.55; 2.04: 5.72; 2.05: 4.7; 2.06: 5.87
| Std Dev. |  | 2.02: 0.32711; 2.03: 0.12111; 2.04: 0.11345; 2.05: 0.31885; 2.06: 0.11916
| Rel.Std.Dev. |  | 2.02: 6.608%; 2.03: 2.671%; 2.04: 1.992%; 2.05: 6.808%; 2.06: 2.03%
| PDM3 |  | 2.02: -3.815%; 2.03: -11.91%; 2.04: 10.66%; 2.05: -8.997%; 2.06: 14.06%
| Z-Score (Absolute) |  | 2.02: 0; 2.03: 0.67; 2.04: 1.21; 2.05: 0.43; 2.06: 1.49
| Analytical results for Rb in OREAS 611 (Certified Value 65 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.20 | 2.21 | 2.23 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 65.6 | 66.5 | 57.4 | 62.7 | 61.7 | 67.3 | 65 | 65 | 62.7 | 64.63 | 68.4581 | 77.3 | 61.5 | 68 | 61.7 | 63.1 | 63 | 70.1 | [5 empty] | 1
| 1 | 2 | 64.8 | 67.4 | 59.6 | 63.1 | 61.1 | 68.3 | 63.3 | 65.2 | 63.2 | 66.24 | 68.3757 | 72.3 | 59.7 | 69 | 64.4 | 66.7 | 62.1 | 70.8 | [5 empty] | 41
| 1 | 3 | 65.9 | 65.3 | 58.4 | 62.9 | 61.6 | 67.9 | 63.9 | 65.3 | 62.4 | 66.96 | 68.6219 | 74.3 | 60.5 | 68.4 | 66.8 | 63.5 | 63.6 | 70.3 | [5 empty] | 16
| 1 | 4 | 64.6 | 66.3 | 60 | 63.2 | 61.9 | 68 | 65.3 | 65.5 | 64.3 | 63.84 | 67.4344 | 68.7 | 60.4 | 68.2 | 64.4 | 63.2 | 62.3 | 68.7 | [5 empty] | 64.5088
| 1 | 5 | 66.7 | 66.6 | 60.5 | 65.4 | 61.9 | 63.2 | 65.1 | 66.4 | 63.8 | 65.54 | 69.3673 | 67.1 | 60.4 | 68 | 65.7 | 62.9 | 62.1 | 68.6 | [5 empty] | 47
| 1 | 6 | 66.4 | 64 | 57.7 | 62.5 | 65.4 | 63.4 | 63.2 | 65.7 | 63.1 | 65.49 | 66.9214 | 82.8 | 61.6 | 68.2 | 64.4 | 62.3 | 62 | 67.8
| Mean |  | 65.6667 | 66.0167 | 58.9333 | 63.3 | 62.2667 | 66.35 | 64.3 | 65.5167 | 63.25 | 65.45 | 68.1965 | 73.75 | 60.6833 | 68.3 | 64.5667 | 63.6167 | 62.5167 | 69.3833
| Median |  | 65.75 | 66.4 | 59 | 63 | 61.8 | 67.6 | 64.45 | 65.4 | 63.15 | 65.515 | 68.4169 | 73.3 | 60.45 | 68.2 | 64.4 | 63.15 | 62.2 | 69.4
| Std Dev. |  | 0.84301 | 1.1957 | 1.2801 | 1.0602 | 1.5629 | 2.3856 | 0.94868 | 0.49565 | 0.70071 | 1.1116 | 0.87852 | 5.773 | 0.73052 | 0.37417 | 1.7072 | 1.5626 | 0.64317 | 1.1788
| Rel.Std.Dev. |  | 1.284% | 1.811% | 2.172% | 1.675% | 2.51% | 3.595% | 1.475% | 0.7565% | 1.108% | 1.698% | 1.288% | 7.828% | 1.204% | 0.5478% | 2.644% | 2.456% | 1.029% | 1.699%
| PDM3 |  | 1.795% | 2.337% | -8.643% | -1.874% | -3.476% | 2.854% | -0.3237% | 1.562% | -1.951% | 1.459% | 5.717% | 14.33% | -5.93% | 5.877% | 0.08969% | -1.383% | -3.088% | 7.556%
| Z-Score (Absolute) |  | 0.26 | 0.39 | 2.36 | 0.66 | 1.07 | 0.52 | 0.28 | 0.2 | 0.68 | 0.17 | 1.24 | 3.4 | 1.68 | 1.28 | 0.17 | 0.54 | 0.97 | 1.7
| Analytical results for Re in OREAS 611 (Indicative Value 2.07 ppb) | [24 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.18 | 2.23 | [7 empty] | ppb
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | < 50 | 3 | 2 | < 2 | < 2 | 2 | 2 | 3 | < 50 | 4 | < 50 | 2 | 3 | < 100 | < 2 | 5 | [7 empty] | 1
| 1 | 2 | < 50 | 3 | 2 | 2 | < 2 | 2 | 2 | < 2 | < 50 | 2 | < 50 | < 2 | 3 | < 100 | < 2 | < 5 | [7 empty] | 14
| 1 | 3 | < 50 | 5 | 2 | < 2 | 3 | < 2 | 3 | 3 | < 50 | 3 | < 50 | 2 | 2 | < 100 | < 2 | < 5 | [7 empty] | 16
| 1 | 4 | < 50 | 7 | 2 | < 2 | 2 | 2 | 2 | 2 | < 50 | 2 | < 50 | 3 | 3 | < 100 | < 2 | 5 | [7 empty] | 2.0741
| 1 | 5 | < 50 | 5 | 3 | < 2 | 3 | < 2 | 2 | 2 | < 50 | 2 | < 50 | 3 | 3 | < 100 | < 2 | 6 | [7 empty] | 11
| 1 | 6 | < 50 | < 2 | 3 | < 2 | 3 | < 2 | 2 | < 2 | < 50 | 2 | < 50 | 2 | 2 | < 100 | < 2 | 5
| Mean |  |  | 4.6 | 2.3333 | 2 | 2.75 | 2 | 2.1667 | 2.5 |  | 2.5 |  | 2.4 | 2.6667 |  |  | 5.25
| Median |  |  | 5 | 2 | 2 | 3 | 2 | 2 | 2.5 |  | 2 |  | 2 | 3 |  |  | 5
| Std Dev. |  |  | 1.6733 | 0.5164 |  | 0.5 | 0 | 0.40825 | 0.57735 |  | 0.83666 |  | 0.54772 | 0.5164 |  |  | 0.5
| Rel.Std.Dev. |  |  | 36.38% | 22.13% |  | 18.18% | 0% | 18.84% | 23.09% |  | 33.47% |  | 22.82% | 19.36% |  |  | 9.524%
| PDM3 |  |  | 121.8% | 12.5% | -3.571% | 32.59% | -3.571% | 4.464% | 20.54% |  | 20.54% |  | 15.71% | 28.57% |  |  | 153.1%
| Z-Score (Absolute) |  | 14.06 | 0.99 | 0.05 | 0.78 | 0.16 | 0.57 | 0.16 | 0.26 | 14.06 | 0.05 | 14.06 | 0.16 | 0.16 | 29.62 | 0.88 | 1.19
| Analytical results for S in OREAS 611 (Certified Value 4.18 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 4.2246 | 4.1 | 4.26 | 4.29 | 4.21 | 4.3 | 4.1 | 4.34 | 4.04 | 4.0957 | 4.3143 | 4.2121 | 3.82 | 4.22 | 4.117 | 4.06 | 4.243 | 4.18 | 3.89 | 4.52 | 4.11 | 4.41 | 4.052 | 1
| 1 | 2 | 4.1162 | 4.04 | 4.17 | 4.31 | 4.17 | 4.23 | 4.05 | 4.38 | 4.1 | 4.2073 | 4.317 | 4.2468 | 3.9 | 4.31 | 4.0962 | 4.06 | 4.099 | 4.15 | 4.04 | 4.44 | 4.12 | 4.44 | 4.059 | 16
| 1 | 3 | 4.1793 | 4.09 | 4.23 | 4.41 | 4.17 | 4.21 | 4.13 | 4.34 | 3.99 | 4.26 | 4.2614 | 4.2755 | 3.8 | 4.36 | 4.093 | 4.11 | 4.151 | 4.06 | 4.06 | 4.52 | 4.09 | 4.42 | 4.066 | 16
| 1 | 4 | 4.132 | 3.92 | 4.42 | 4.36 | 4.2 | 4.24 | 4.14 | 4.36 | 4.04 | 4.1023 | 4.2551 | 4.2513 | 3.89 | 4.34 | 4.0718 | 4 | 4.184 | 4.12 | 4.07 | 4.62 | 4.1 | 4.42 | 4.073 | 4.177
| 1 | 5 | 4.156 | 4.04 | 4.44 | 4.47 | 4.17 | 3.9 | 4.08 | 4.52 | 4.08 | 4.1928 | 4.3433 | 4.2522 | 3.92 | 4.23 | 4.0707 | 4.04 | 4.099 | 4.02 | 3.95 | 4.52 | 4.12 | 4.46 | 4.0704 | 48
| 1 | 6 | 4.2221 | 3.72 | 4.29 | 4.33 | 4.07 | 4.02 | 4.06 | 4.44 | 4.03 | 4.1196 | 4.293 | 4.4462 | 3.8 | 4.35 | 4.1196 | 4.01 | 4.065 | 4.07 | 4.03 | 4.37 | 4.1 | 4.4 | 4.059
| Mean |  | 4.1717 | 3.985 | 4.3017 | 4.3617 | 4.165 | 4.15 | 4.0933 | 4.3967 | 4.0467 | 4.163 | 4.2974 | 4.2807 | 3.855 | 4.3017 | 4.0947 | 4.0467 | 4.1402 | 4.1 | 4.0067 | 4.4983 | 4.1067 | 4.425 | 4.0632
| Median |  | 4.1677 | 4.04 | 4.275 | 4.345 | 4.17 | 4.22 | 4.09 | 4.37 | 4.04 | 4.1562 | 4.3037 | 4.2518 | 3.855 | 4.325 | 4.0946 | 4.05 | 4.125 | 4.095 | 4.035 | 4.52 | 4.105 | 4.42 | 4.0625
| Std Dev. |  | 0.045385 | 0.14474 | 0.10722 | 0.067651 | 0.049699 | 0.15492 | 0.036697 | 0.070899 | 0.038816 | 0.066868 | 0.034293 | 0.083605 | 0.054314 | 0.061779 | 0.021093 | 0.039833 | 0.065807 | 0.060332 | 0.071181 | 0.084951 | 0.012111 | 0.021679 | 0.0079553
| Rel.Std.Dev. |  | 1.088% | 3.632% | 2.493% | 1.551% | 1.193% | 3.733% | 0.8965% | 1.613% | 0.9592% | 1.606% | 0.798% | 1.953% | 1.409% | 1.436% | 0.5151% | 0.9843% | 1.589% | 1.472% | 1.777% | 1.888% | 0.2949% | 0.4899% | 0.1958%
| PDM3 |  | -0.1266% | -4.596% | 2.985% | 4.421% | -0.287% | -0.6461% | -2.003% | 5.259% | -3.12% | -0.3361% | 2.882% | 2.483% | -7.709% | 2.985% | -1.97% | -3.12% | -0.8815% | -1.843% | -4.078% | 7.693% | -1.684% | 5.938% | -2.723%
| Z-Score (Absolute) |  | 0.14 | 1.08 | 0.99 | 1.38 | 0.1 | 0 | 0.37 | 1.61 | 0.67 | 0.08 | 0.96 | 0.85 | 1.93 | 0.99 | 0.36 | 0.67 | 0.06 | 0.33 | 0.94 | 2.27 | 0.28 | 1.79 | 0.57
| Analytical results for Sb in OREAS 611 (Certified Value 365 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 322.3 | 365.71 | 41.9 | 372 | 356 | 376 | 368 | 381 | 369 | 363.68 | 327.8801 | 363.05 | 408 | 428 | 417.4373 | 327.8 | 314.498 | 364.7 | 350 | 374 | 307.2 | 340 | 355.7 | 1
| 1 | 2 | 322 | 378.1 | 42.3 | 379 | 352 | 380 | 364 | 386 | 372 | 366.83 | 351.6556 | 379.01 | 401 | 437 | 420.7134 | 333.5 | 325.018 | 364.9 | 346 | 365 | 311.2 | 340 | 359.4 | 43
| 1 | 3 | 328.7 | 377.64 | 70 | 390 | 359 | 374 | 370 | 386 | 370 | 368.38 | 322.6401 | 384.25 | 402 | 433 | 414.0459 | 339.6 | 311.386 | 361 | 345 | 370 | 314.9 | 340 | 357.1 | 16
| 1 | 4 | 326.5 | 383.76 | 70.2 | 387 | 359 | 374 | 374 | 382 | 373 | 360.26 | 346.5787 | 383.01 | 404 | 438 | 395.4615 | 341.2 | 295.659 | 351.2 | 348 | 385 | 316 | 350 | 354.8 | 364.7518
| 1 | 5 | 333.8 | 379.19 | 69.2 | 393 | 352 | 348 | 372 | 401 | 370 | 363.36 | 338.1144 | 383.61 | 403 | 438 | 416.2266 | 360.8 | 330.724 | 363 | 345 | 371 | 306.8 | 340 | 360.5 | 49
| 1 | 6 | 329.7 | 355.4 | 30.2 | 385 | 380 | 362 | 362 | 395 | 372 | 368.88 | 360.9115 | 382.87 | 399 | 434 | 405.5604 | 339.2 | 328.299 | 356.2 | 345 | 369 | 302.3 | 340 | 355.2
| Mean |  | 327.1667 | 373.3 | 53.9667 | 384.3333 | 359.6667 | 369 | 368.3333 | 388.5 | 371 | 365.2317 | 341.2967 | 379.3 | 402.8333 | 434.6667 | 411.5742 | 340.35 | 317.5973 | 360.1667 | 346.5 | 372.3333 | 309.7333 | 341.6667 | 357.1167
| Median |  | 327.6 | 377.87 | 55.75 | 386 | 357.5 | 374 | 369 | 386 | 371 | 365.255 | 342.3465 | 382.94 | 402.5 | 435.5 | 415.1362 | 339.4 | 319.758 | 362 | 345.5 | 370.5 | 309.2 | 340 | 356.4
| Std Dev. |  | 4.5518 | 10.6207 | 17.8841 | 7.6855 | 10.4435 | 11.9164 | 4.6332 | 7.8677 | 1.5492 | 3.3587 | 14.5475 | 8.1699 | 3.0605 | 3.8816 | 9.3962 | 11.1824 | 13.2014 | 5.4372 | 2.0736 | 6.8605 | 5.2611 | 4.0825 | 2.3541
| Rel.Std.Dev. |  | 1.391% | 2.845% | 33.14% | 2% | 2.904% | 3.229% | 1.258% | 2.025% | 0.4176% | 0.9196% | 4.262% | 2.154% | 0.7597% | 0.893% | 2.283% | 3.286% | 4.157% | 1.51% | 0.5985% | 1.843% | 1.699% | 1.195% | 0.6592%
| PDM3 |  | -10.3% | 2.344% | -85.2% | 5.368% | -1.394% | 1.165% | 0.9819% | 6.511% | 1.713% | 0.1316% | -6.43% | 3.989% | 10.44% | 19.17% | 12.84% | -6.69% | -12.93% | -1.257% | -5.004% | 2.079% | -15.08% | -6.329% | -2.093%
| Z-Score (Absolute) |  | 1.34 | 0.28 | 10.99 | 0.67 | 0.2 | 0.13 | 0.11 | 0.82 | 0.2 | 0 | 0.84 | 0.5 | 1.33 | 2.45 | 1.64 | 0.88 | 1.68 | 0.18 | 0.66 | 0.25 | 1.96 | 0.83 | 0.29
| Analytical results for Sc in OREAS 611 (Certified Value 3.2 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3 | 3.3 | 4 | 3.2 | 3.2 | 3.4 | 3.1 | 3.2 | 3 | 3.5 | 3.0367 | 3.1 | 3.6 | 4 | < 5 | 2.5 | 3.053 | 2.8 | 2.9 | 3.1 | 3.8 |  |  | 1
| 1 | 2 | 3 | 3.4 | 4 | 3.2 | 3.1 | 3.4 | 3 | 3.2 | 3 | 3.3 | 2.9814 | 3 | 3.6 | 4 | < 5 | 3.1 | 2.987 | 2.9 | 2.9 | 3 | 3.7 |  |  | 44
| 1 | 3 | 3 | 3.5 | 4 | 3.1 | 3.2 | 3.3 | 3.1 | 3.1 | 3 | 3.5 | 3.4076 | 3 | 3.6 | 4 | < 5 | 3.2 | 3.037 | 2.8 | 2.9 | 3 | 3.4 |  |  | 16
| 1 | 4 | 3 | 3.4 | 4 | 3.1 | 3.2 | 3.4 | 3.3 | 3.2 | 3 | 3.4 | 3.204 | 3 | 3.5 | 4 | < 5 | 3.2 | 3.072 | 2.9 | 2.9 | 3.1 | 3.3 |  |  | 3.2
| 1 | 5 | 3 | 3.3 | 4 | 3.5 | 3.3 | 3.2 | 3.2 | 3.3 | 3 | 3.3 | 3.3644 | 3.1 | 3.5 | 4 | < 5 | 3.2 | 2.833 | 2.9 | 2.9 | 3 | 3.5 |  |  | 50
| 1 | 6 | 3 | 3.2 | 4 | 3.4 | 3.3 | 3.2 | 3.2 | 3.3 | 3 | 3.5 | 3.3106 | 3.2 | 3.6 | 4 | < 5 | 3.1 | 2.95 | 2.7 | 2.9 | 3 | 3.4
| Mean |  | 3 | 3.35 | 4 | 3.25 | 3.2167 | 3.3167 | 3.15 | 3.2167 | 3 | 3.4167 | 3.2174 | 3.0667 | 3.5667 | 4 |  | 3.05 | 2.9887 | 2.8333 | 2.9 | 3.0333 | 3.5167
| Median |  | 3 | 3.35 | 4 | 3.2 | 3.2 | 3.35 | 3.15 | 3.2 | 3 | 3.45 | 3.2573 | 3.05 | 3.6 | 4 |  | 3.15 | 3.012 | 2.85 | 2.9 | 3 | 3.45
| Std Dev. |  | 0 | 0.10488 | 0 | 0.16432 | 0.075277 | 0.098319 | 0.10488 | 0.075277 | 0 | 0.098319 | 0.17607 | 0.08165 | 0.05164 | 0 |  | 0.27386 | 0.088487 | 0.08165 | 0 | 0.05164 | 0.19408
| Rel.Std.Dev. |  | 0% | 3.131% | 0% | 5.056% | 2.34% | 2.964% | 3.33% | 2.34% | 0% | 2.878% | 5.472% | 2.662% | 1.448% | 0% |  | 8.979% | 2.961% | 2.882% | 0% | 1.702% | 5.519%
| PDM3 |  | -6.249% | 4.689% | 25% | 1.564% | 0.5219% | 3.647% | -1.561% | 0.5219% | -6.249% | 6.772% | 0.546% | -4.166% | 11.46% | 25% |  | -4.686% | -6.603% | -11.46% | -9.374% | -5.207% | 9.897%
| Z-Score (Absolute) |  | NA | 0.54 | NA | 0.13 | 0 | 0.4 | 0.27 | 0 | NA | 0.81 | 0 | 0.61 | 1.42 | NA | 2.9 | 0.67 | 0.92 | 1.55 | 1.28 | 0.74 | 1.21
Results from laboratories 2, 4, 10 & 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 611 (Certified Value 34.8 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 33 | 35 | 36.9 | 37 | 36 | 33 | 36 | 39 | 29 | 31.8 | 35.6309 | 30.7 | 40.3 | 30 | 42.4253 | 23.3 | 35 | 29 | [5 empty] | 1
| 1 | 2 | 33 | 35.9 | 38.4 | 35 | 34 | 33 | 36 | 37 | 30 | 34.1 | 38.0071 | 31.7 | 39.1 | 30 | 35.6556 | 24.1 | 38 | 28 | [5 empty] | 45
| 1 | 3 | 34 | 35.3 | 37.7 | 33 | 35 | 34 | 35 | 37 | 29 | 34.4 | 37.8864 | 34.4 | 39.5 | 35 | 32.3669 | 23.5 | 37 | 29 | [5 empty] | 16
| 1 | 4 | 32 | 35.8 | 39 | 36 | 36 | 33 | 38 | 38 | 30 | 32.8 | 36.7799 | 29.9 | 39.6 | 35 | 38.309 | 23.6 | 35 | 28 | [5 empty] | 34.7806
| 1 | 5 | 33 | 36.9 | 39.8 | 36 | 33 | 32 | 35 | 39 | 30 | 32.5 | 36.892 | 30.3 | 38 | 35 | 37.3113 | 25.2 | 36 | 27 | [5 empty] | 51
| 1 | 6 | 34 | 34.7 | 40.3 | 35 | 38 | 33 | 37 | 40 | 30 | 33 | 36.2112 | 31.5 | 38.4 | 35 | 42.4411 | 22.7 | 35 | 27
| Mean |  | 33.1667 | 35.6 | 38.6833 | 35.3333 | 35.3333 | 33 | 36.1667 | 38.3333 | 29.6667 | 33.1 | 36.9012 | 31.4167 | 39.15 | 33.3333 | 38.0849 | 23.7333 | 36 | 28
| Median |  | 33 | 35.55 | 38.7 | 35.5 | 35.5 | 33 | 36 | 38.5 | 30 | 32.9 | 36.8359 | 31.1 | 39.3 | 35 | 37.8101 | 23.55 | 35.5 | 28
| Std Dev. |  | 0.75277 | 0.78486 | 1.2797 | 1.3663 | 1.7512 | 0.63246 | 1.169 | 1.2111 | 0.5164 | 0.98387 | 0.92718 | 1.6154 | 0.84083 | 2.582 | 3.9258 | 0.8501 | 1.2649 | 0.89443
| Rel.Std.Dev. |  | 2.27% | 2.205% | 3.308% | 3.867% | 4.956% | 1.917% | 3.232% | 3.159% | 1.741% | 2.972% | 2.513% | 5.142% | 2.148% | 7.746% | 10.31% | 3.582% | 3.514% | 3.194%
| PDM3 |  | -4.64% | 2.356% | 11.22% | 1.589% | 1.589% | -5.119% | 3.985% | 10.21% | -14.7% | -4.832% | 6.097% | -9.672% | 12.56% | -4.161% | 9.5% | -31.76% | 3.506% | -19.5%
| Z-Score (Absolute) |  | 0.64 | 0.08 | 0.99 | 0 | 0 | 0.69 | 0.25 | 0.89 | 1.67 | 0.66 | 0.46 | 1.16 | 1.13 | 0.59 | 0.81 | 3.43 | 0.2 | 2.17
| Analytical results for Sm in OREAS 611 (Certified Value 3.8 ppm) | [24 empty] | Certified
Sm | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab | [5 empty] | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 | [5 empty] | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS | [5 empty] | 2
3
| 1 | 1 | 2.02: 3.3; 2.03: 2.9; 2.04: 4.04; 2.05: 3.8; 2.06: 4.05 | [5 empty] | 1
| 1 | 2 | 2.02: 3.5; 2.03: 3.7; 2.04: 4.05; 2.05: 3.5; 2.06: 4.05 | [5 empty] | 8
| 1 | 3 | 2.02: 3.9; 2.03: 3.7; 2.04: 3.95; 2.05: 3.8; 2.06: 4 | [5 empty] | 16
| 1 | 4 | 2.02: 3.8; 2.03: 3.7; 2.04: 3.88; 2.05: 3.6; 2.06: 4.05 | [5 empty] | 3.8013
| 1 | 5 | 2.02: 3.8; 2.03: 3.8; 2.04: 4.02; 2.05: 3.5; 2.06: 4.1 | [5 empty] | 52
| 1 | 6 | 2.02: 3.5; 2.03: 3.5; 2.04: 4.07; 2.05: 3.7; 2.06: 4
| Mean |  | 2.02: 3.6333; 2.03: 3.55; 2.04: 4.0017; 2.05: 3.65; 2.06: 4.0417
| Median |  | 2.02: 3.65; 2.03: 3.7; 2.04: 4.03; 2.05: 3.65; 2.06: 4.05
| Std Dev. |  | 2.02: 0.23381; 2.03: 0.33317; 2.04: 0.072503; 2.05: 0.13784; 2.06: 0.037639
| Rel.Std.Dev. |  | 2.02: 6.435%; 2.03: 9.385%; 2.04: 1.812%; 2.05: 3.776%; 2.06: 0.9313%
| PDM3 |  | 2.02: -4.42%; 2.03: -6.612%; 2.04: 5.27%; 2.05: -3.981%; 2.06: 6.322%
| Z-Score (Absolute) |  | 2.02: 0.11; 2.03: 0.67; 2.04: 2.37; 2.05: 0; 2.06: 2.64
| Analytical results for Sn in OREAS 611 (Certified Value 32.1 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 32.1 | 34.7 | 30 | 31.5 | 31.9 | 32.7 | 33 | 34.9 | 29.3 | 32.7 | 35.3308 | 33.7 | 35.1 | 35 | 28.7865 | 29.5 | 31.626 | 31 | 30.4 | 29 | 35.3 | 30 | 27.6 | 1
| 1 | 2 | 31.5 | 34.1 | 30 | 31.8 | 31.7 | 32.2 | 31.6 | 34.9 | 29.3 | 32.9 | 34.6244 | 34.1 | 34.3 | 35 | 31.7425 | 29.4 | 32.085 | 31 | 29.4 | 27 | 35.7 | 31 | 27.66 | 18
| 1 | 3 | 32.3 | 35.3 | 30 | 32.2 | 31.8 | 32.3 | 32.2 | 34.3 | 30.3 | 33.4 | 33.2855 | 37.1 | 34 | 35 | 32.3878 | 29.4 | 31.231 | 30.5 | 29.8 | 28 | 35.6 | 30 | 26.72 | 16
| 1 | 4 | 32.4 | 34.5 | 31 | 32.2 | 31.8 | 32 | 33.7 | 34 | 30 | 32.4 | 34.8287 | 37.5 | 34.6 | 35 | 30.4266 | 29.3 | 31.679 | 30.1 | 29.6 | 30 | 35.9 | 31 | 27.78 | 32.0502
| 1 | 5 | 32.7 | 35.3 | 32 | 32.3 | 32 | 31.2 | 33 | 35.9 | 29.6 | 32.9 | 35.1833 | 35.1 | 34.5 | 34 | 31.7924 | 31.1 | 31.427 | 30.8 | 29.8 | 29 | 35 | 31 | 27.84 | 53
| 1 | 6 | 32.1 | 32.7 | 32 | 31.9 | 33 | 31.3 | 32.1 | 36.7 | 30.2 | 32.9 | 33.8117 | 36.4 | 34.8 | 34 | 31.0046 | 30.1 | 31.264 | 30.2 | 29.6 | 26 | 34.2 | 30 | 27.9
| Mean |  | 32.1833 | 34.4333 | 30.8333 | 31.9833 | 32.0333 | 31.95 | 32.6 | 35.1167 | 29.7833 | 32.8667 | 34.5107 | 35.65 | 34.55 | 34.6667 | 31.0234 | 29.8 | 31.552 | 30.6 | 29.7667 | 28.1667 | 35.2833 | 30.5 | 27.5833
| Median |  | 32.2 | 34.6 | 30.5 | 32.05 | 31.85 | 32.1 | 32.6 | 34.9 | 29.8 | 32.9 | 34.7265 | 35.75 | 34.55 | 35 | 31.3735 | 29.45 | 31.5265 | 30.65 | 29.7 | 28.5 | 35.45 | 30.5 | 27.72
| Std Dev. |  | 0.40208 | 0.96885 | 0.98319 | 0.30605 | 0.48442 | 0.58907 | 0.76681 | 1.0128 | 0.4446 | 0.3266 | 0.8036 | 1.5871 | 0.38341 | 0.5164 | 1.2908 | 0.69857 | 0.31853 | 0.39497 | 0.34448 | 1.472 | 0.61779 | 0.54772 | 0.43734
| Rel.Std.Dev. |  | 1.249% | 2.814% | 3.189% | 0.9569% | 1.512% | 1.844% | 2.352% | 2.884% | 1.493% | 0.9937% | 2.329% | 4.452% | 1.11% | 1.49% | 4.161% | 2.344% | 1.01% | 1.291% | 1.157% | 5.226% | 1.751% | 1.796% | 1.586%
| PDM3 |  | 0.4153% | 7.436% | -3.797% | -0.2087% | -0.05274% | -0.3128% | 1.715% | 9.568% | -7.073% | 2.547% | 7.677% | 11.23% | 7.8% | 8.164% | -3.204% | -7.021% | -1.555% | -4.525% | -7.125% | -12.12% | 10.09% | -4.837% | -13.94%
| Z-Score (Absolute) |  | 0.06 | 0.76 | 0.36 | 0 | 0.02 | 0.01 | 0.19 | 0.97 | 0.68 | 0.27 | 0.78 | 1.13 | 0.79 | 0.83 | 0.3 | 0.67 | 0.13 | 0.43 | 0.68 | 1.18 | 1.02 | 0.46 | 1.36
| Analytical results for Sr in OREAS 611 (Certified Value 317 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 311.2 | 268 | 190 | 319 | 322 | 337 | 313 | 326 | 330 | 328.28 | 313.3061 | 239 | 311 | 339 | 348.6979 | 259 | 296.9 | 268.8 | 327 | 304 | 300 | 230 | 346.4 | 1
| 1 | 2 | 308.2 | 285 | 195 | 320 | 316 | 330 | 305 | 330 | 313 | 332.67 | 317.0538 | 262 | 321 | 341 | 348.3982 | 249.2 | 299.912 | 272.3 | 325 | 305 | 292 | 230 | 342 | 19
| 1 | 3 | 320.3 | 271 | 181 | 333 | 308 | 328 | 326 | 328 | 313 | 330.93 | 327.5237 | 270 | 314 | 332 | 349.551 | 272.8 | 306.963 | 263.7 | 328 | 322 | 289 | 230 | 344.6 | 16
| 1 | 4 | 314.2 | 287 | 185 | 326 | 321 | 332 | 330 | 326 | 329 | 316.87 | 315.5638 | 249 | 327 | 330 | 347.3768 | 244.6 | 303.183 | 275.4 | 314 | 313 | 278 | 230 | 343.2 | 316.8633
| 1 | 5 | 314.3 | 274 | 196 | 335 | 317 | 321 | 321 | 342 | 326 | 334.23 | 324.7011 | 246 | 317 | 339 | 350.2206 | 282.9 | 288.103 | 269.4 | 331 | 305 | 279 | 230 | 345.8 | 54
| 1 | 6 | 316.5 | 271 | 161 | 323 | 317 | 326 | 314 | 332 | 325 | 328.02 | 323.5069 | 275 | 315 | 334 | 350.2545 | 271 | 291.049 | 275.1 | 316 | 325 | 284 | 230 | 344.4
| Mean |  | 314.1167 | 276 | 184.6667 | 326 | 316.8333 | 329 | 318.1667 | 330.6667 | 322.6667 | 328.5 | 320.2759 | 256.8333 | 317.5 | 335.8333 | 349.0832 | 263.25 | 297.685 | 270.7833 | 323.5 | 312.3333 | 287 | 230 | 344.4
| Median |  | 314.25 | 272.5 | 187.5 | 324.5 | 317 | 329 | 317.5 | 329 | 325.5 | 329.605 | 320.2803 | 255.5 | 316 | 336.5 | 349.1244 | 265 | 298.406 | 270.85 | 326 | 309 | 286.5 | 230 | 344.5
| Std Dev. |  | 4.1825 | 8 | 12.9409 | 6.6933 | 4.9565 | 5.4406 | 9.2394 | 6.0222 | 7.7115 | 6.1913 | 5.7222 | 14.3306 | 5.7184 | 4.446 | 1.1326 | 14.8369 | 7.1796 | 4.4332 | 6.892 | 9.288 | 8.3905 | 0 | 1.6248
| Rel.Std.Dev. |  | 1.332% | 2.899% | 7.008% | 2.053% | 1.564% | 1.654% | 2.904% | 1.821% | 2.39% | 1.885% | 1.787% | 5.58% | 1.801% | 1.324% | 0.3245% | 5.636% | 2.412% | 1.637% | 2.13% | 2.974% | 2.924% | 0% | 0.4718%
| PDM3 |  | -0.8668% | -12.9% | -41.72% | 2.883% | -0.009469% | 3.83% | 0.4113% | 4.356% | 1.831% | 3.672% | 1.077% | -18.95% | 0.2009% | 5.987% | 10.17% | -16.92% | -6.053% | -14.54% | 2.094% | -1.43% | -9.425% | -27.41% | 8.69%
| Z-Score (Absolute) |  | 0.17 | 2.13 | 6.8 | 0.44 | 0.03 | 0.59 | 0.03 | 0.67 | 0.26 | 0.56 | 0.14 | 3.11 | 0 | 0.94 | 1.62 | 2.78 | 1.01 | 2.39 | 0.31 | 0.26 | 1.56 | 4.48 | 1.38
| Analytical results for Ta in OREAS 611 (Certified Value 0.71 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.65 | 0.7 | 0.2 | 0.74 | 0.66 | 0.74 | 0.66 | 0.67 | 0.7 | 0.69 | 0.50058 | 0.8 | 0.84 | 0.8 | < 10 | 0.66 | 0.76 | 0.79 | 0.8 |  |  |  |  | 1
| 1 | 2 | 0.67 | 0.7 | 0.2 | 0.69 | 0.66 | 0.7 | 0.64 | 0.67 | 0.72 | 0.74 | 0.50654 | 0.8 | 0.76 | 0.8 | < 10 | 0.73 | 0.79 | 0.75 | 0.9 |  |  |  |  | 9
| 1 | 3 | 0.67 | 0.7 | 0.2 | 0.73 | 0.64 | 0.7 | 0.68 | 0.68 | 0.68 | 0.74 | 0.50698 | 0.8 | 0.78 | 0.8 | < 10 | 0.73 | 0.74 | 0.66 | 0.7 |  |  |  |  | 16
| 1 | 4 | 0.68 | 0.7 | 0.2 | 0.72 | 0.66 | 0.72 | 0.7 | 0.65 | 0.73 | 0.73 | 0.49729 | 0.8 | 0.78 | 0.8 | < 10 | 0.8 | 0.74 | 0.77 | 0.8 |  |  |  |  | 0.71394
| 1 | 5 | 0.71 | 0.7 | < 0.1 | 0.75 | 0.65 | 0.67 | 0.65 | 0.69 | 0.7 | 0.73 | 0.50274 | 0.7 | 0.77 | 0.8 | < 10 | 0.66 | 0.78 | 0.78 | 0.8 |  |  |  |  | 55
| 1 | 6 | 0.69 | 0.7 | < 0.1 | 0.72 | 0.66 | 0.67 | 0.68 | 0.68 | 0.7 | 0.72 | 0.5003 | 0.8 | 0.78 | 0.8 | < 10 | 0.73 | 0.75 | 0.8 | 0.7
| Mean |  | 0.67833 | 0.7 | 0.2 | 0.725 | 0.655 | 0.7 | 0.66833 | 0.67333 | 0.705 | 0.725 | 0.5024 | 0.78333 | 0.785 | 0.8 |  | 0.71833 | 0.76 | 0.75833 | 0.78333
| Median |  | 0.675 | 0.7 | 0.2 | 0.725 | 0.66 | 0.7 | 0.67 | 0.675 | 0.7 | 0.73 | 0.50166 | 0.8 | 0.78 | 0.8 |  | 0.73 | 0.755 | 0.775 | 0.8
| Std Dev. |  | 0.020412 | 1.2162e-16 | 0 | 0.020736 | 0.0083666 | 0.027568 | 0.022286 | 0.013663 | 0.017607 | 0.018708 | 0.0037952 | 0.040825 | 0.028107 | 1.2162e-16 |  | 0.052694 | 0.020976 | 0.051153 | 0.075277
| Rel.Std.Dev. |  | 3.009% | 1.737e-14% | 0% | 2.86% | 1.277% | 3.938% | 3.335% | 2.029% | 2.497% | 2.58% | 0.7554% | 5.212% | 3.581% | 1.52e-14% |  | 7.336% | 2.76% | 6.745% | 9.61%
| PDM3 |  | -4.988% | -1.953% | -71.99% | 1.549% | -8.256% | -1.953% | -6.389% | -5.688% | -1.253% | 1.549% | -29.63% | 9.719% | 9.953% | 12.05% |  | 0.6147% | 6.451% | 6.217% | 9.719%
| Z-Score (Absolute) |  | 0.55 | NA | NA | 0.22 | 0.94 | 0.19 | 0.72 | 0.63 | 0.11 | 0.22 | 3.46 | NA | 1.21 | NA | 70.82 | 0.11 | 0.8 | 0.77 | NA
Results from laboratories 3, 4, 14, 16 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 611 (Certified Value 0.31 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16; 2.07: 2.18; 2.08: 2.20; 2.09: 2.21 |  |  |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.3; 2.06: 0.4; 2.07: 0.2; 2.08: 0.26; 2.09: 0.35 |  |  |  |  | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.25; 2.09: 0.35 |  |  |  |  | 10
| 1 | 3 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.38; 2.05: 0.3; 2.06: 0.38; 2.07: 0.3; 2.08: 0.25; 2.09: 0.34 |  |  |  |  | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.3; 2.04: 0.37; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.26; 2.09: 0.35 |  |  |  |  | 0.31079
| 1 | 5 | 2.02: 0.2; 2.03: 0.4; 2.04: 0.38; 2.05: 0.3; 2.06: 0.38; 2.07: 0.3; 2.08: 0.26; 2.09: 0.32 |  |  |  |  | 56
| 1 | 6 | 2.02: 0.2; 2.03: 0.4; 2.04: 0.38; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.24; 2.09: 0.35
| Mean |  | 2.02: 0.2; 2.03: 0.33333; 2.04: 0.375; 2.05: 0.3; 2.06: 0.39333; 2.07: 0.28333; 2.08: 0.25333; 2.09: 0.34333
| Median |  | 2.02: 0.2; 2.03: 0.3; 2.04: 0.375; 2.05: 0.3; 2.06: 0.4; 2.07: 0.3; 2.08: 0.255; 2.09: 0.35
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 0.05164; 2.04: 0.0054772; 2.05: 0; 2.06: 0.010328; 2.07: 0.040825; 2.08: 0.008165; 2.09: 0.012111
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 15.49%; 2.04: 1.461%; 2.05: 0%; 2.06: 2.626%; 2.07: 14.41%; 2.08: 3.223%; 2.09: 3.527%
| PDM3 |  | 2.02: -35.65%; 2.03: 7.253%; 2.04: 20.66%; 2.05: -3.472%; 2.06: 26.56%; 2.07: -8.835%; 2.08: -18.49%; 2.09: 10.47%
| Z-Score (Absolute) |  | 2.02: 1.72; 2.03: 0.25; 2.04: 0.86; 2.05: 0.25; 2.06: 1.13; 2.07: 0.49; 2.08: 0.93; 2.09: 0.39
| Analytical results for Te in OREAS 611 (Certified Value 49.3 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 44.7 | 50.45 | 28.8 | 50 | 49.3 | 48.6 | 50.1 | 54.2 | 47.6 | 49.5 | 53.2064 | 45.18 | 50.1 | 35.6 | 51.3442 | 43.06 | 51.6 | 47.58 | 44.3 |  |  |  |  | 1
| 1 | 2 | 44 | 51.99 | 28.8 | 49.3 | 49 | 48.8 | 48.8 | 53.7 | 46.7 | 49.9 | 50.907 | 45.53 | 48.9 | 35.6 | 58.4854 | 47.55 | 50.2 | 48.02 | 45.8 |  |  |  |  | 34
| 1 | 3 | 45.1 | 50.93 | 32.3 | 48.9 | 50 | 48.3 | 49.4 | 53.5 | 46.5 | 51.4 | 52.3869 | 50.61 | 49.7 | 35.8 | 52.5908 | 47.74 | 49.9 | 47.84 | 46.9 |  |  |  |  | 16
| 1 | 4 | 44.6 | 51.6 | 32 | 49.8 | 49.4 | 47.6 | 51.6 | 53.8 | 48.6 | 48.1 | 49.5995 | 49.87 | 49.1 | 36.2 | 56.3264 | 48.2 | 48.4 | 47.15 | 46.2 |  |  |  |  | 49.2816
| 1 | 5 | 44.3 | 49.95 | 31.5 | 55.2 | 49.4 | 47.5 | 49.4 | 55.4 | 47 | 49.5 | 50.8327 | 47.24 | 49.1 | 36 | 54.5097 | 48.23 | 48.3 | 46.59 | 46.4 |  |  |  |  | 57
| 1 | 6 | 45 | 46.94 | 29.4 | 52.5 | 50.2 | 48.2 | 51.8 | 55.1 | 46.6 | 50.3 | 52.265 | 49.12 | 49.6 | 35.6 | 54.8252 | 52.3 | 49.4 | 47.32 | 43.7
| Mean |  | 44.6167 | 50.31 | 30.4667 | 50.95 | 49.55 | 48.1667 | 50.1833 | 54.2833 | 47.1667 | 49.7833 | 51.5329 | 47.925 | 49.4167 | 35.8 | 54.6803 | 47.8467 | 49.6333 | 47.4167 | 45.55
| Median |  | 44.65 | 50.69 | 30.45 | 49.9 | 49.4 | 48.25 | 49.75 | 54 | 46.85 | 49.7 | 51.586 | 48.18 | 49.35 | 35.7 | 54.6675 | 47.97 | 49.65 | 47.45 | 46
| Std Dev. |  | 0.41673 | 1.8098 | 1.6415 | 2.4337 | 0.45497 | 0.52409 | 1.2465 | 0.78846 | 0.80664 | 1.0852 | 1.318 | 2.2874 | 0.45789 | 0.25298 | 2.5587 | 2.9365 | 1.234 | 0.51648 | 1.2661
| Rel.Std.Dev. |  | 0.934% | 3.597% | 5.388% | 4.777% | 0.9182% | 1.088% | 2.484% | 1.452% | 1.71% | 2.18% | 2.558% | 4.773% | 0.9266% | 0.7067% | 4.679% | 6.137% | 2.486% | 1.089% | 2.78%
| PDM3 |  | -9.466% | 2.087% | -38.18% | 3.385% | 0.5445% | -2.262% | 1.83% | 10.15% | -4.292% | 1.018% | 4.568% | -2.753% | 0.274% | -27.36% | 10.95% | -2.912% | 0.7136% | -3.784% | -7.572%
| Z-Score (Absolute) |  | 2.06 | 0.38 | 8.14 | 0.66 | 0.06 | 0.54 | 0.33 | 2.09 | 0.97 | 0.16 | 0.91 | 0.64 | 0 | 5.85 | 2.26 | 0.67 | 0.09 | 0.86 | 1.66
| Analytical results for Th in OREAS 611 (Certified Value 8.66 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 9.14 | 6.6 | 3.2 | 8.66 | 8.41 | 9.11 | 8.25 | 10.45 | 8.41 | 9.12 | 8.5472 | 6.2 | 8.4 | 9.6 | < 10 | 7.17 | 6 | 8.6 | 10.4 |  |  |  |  | 1
| 1 | 2 | 9.11 | 6.6 | 3.3 | 9.04 | 8.56 | 8.75 | 8.1 | 10.6 | 8.13 | 9.26 | 8.975 | 6.6 | 8.9 | 9.7 | < 10 | 9.11 | 5.6 | 7.9 | 10.3 |  |  |  |  | 35
| 1 | 3 | 9.47 | 6.1 | 1.9 | 8.95 | 8.55 | 8.96 | 8.37 | 10.8 | 8.11 | 9.29 | 8.3355 | 7.1 | 8.5 | 9.2 | < 10 | 8.36 | 5.5 | 7.1 | 10.2 |  |  |  |  | 16
| 1 | 4 | 9.56 | 7.3 | 1.7 | 8.55 | 8.51 | 8.91 | 9.06 | 10.3 | 8.1 | 8.96 | 8.6688 | 6.5 | 9.2 | 9.2 | 12.4028 | 7.74 | 5.8 | 8 | 10.2 |  |  |  |  | 8.6627
| 1 | 5 | 9.21 | 7.1 | 1.7 | 9.6 | 8.52 | 8.69 | 8.71 | 11.1 | 8.12 | 9.6 | 8.703 | 6 | 9.2 | 9.2 | < 10 | 8.93 | 5.9 | 8.1 | 9.6 |  |  |  |  | 58
| 1 | 6 | 9.94 | 6.8 | 2.7 | 8.8 | 9.19 | 8.61 | 8.24 | 10.5 | 8.17 | 9.3 | 8.8946 | 7.1 | 8.6 | 9.1 | 10.7977 | 7.8 | 5.6 | 8.4 | 9.9
| Mean |  | 9.405 | 6.75 | 2.4167 | 8.9333 | 8.6233 | 8.8383 | 8.455 | 10.625 | 8.1733 | 9.255 | 8.6873 | 6.5833 | 8.8 | 9.3333 | 11.6002 | 8.185 | 5.7333 | 8.0167 | 10.1
| Median |  | 9.34 | 6.7 | 2.3 | 8.875 | 8.535 | 8.83 | 8.31 | 10.55 | 8.125 | 9.275 | 8.6859 | 6.55 | 8.75 | 9.2 | 11.6002 | 8.08 | 5.7 | 8.05 | 10.2
| Std Dev. |  | 0.3193 | 0.42308 | 0.74409 | 0.37297 | 0.28268 | 0.18745 | 0.36126 | 0.28592 | 0.11843 | 0.21333 | 0.23225 | 0.4535 | 0.35214 | 0.25033 | 1.135 | 0.7507 | 0.19664 | 0.51929 | 0.29665
| Rel.Std.Dev. |  | 3.395% | 6.268% | 30.79% | 4.175% | 3.278% | 2.121% | 4.273% | 2.691% | 1.449% | 2.305% | 2.673% | 6.889% | 4.002% | 2.682% | 9.784% | 9.172% | 3.43% | 6.478% | 2.937%
| PDM3 |  | 8.569% | -22.08% | -72.1% | 3.124% | -0.4545% | 2.027% | -2.398% | 22.65% | -5.649% | 6.837% | 0.2842% | -24% | 1.585% | 7.742% | 33.91% | -5.515% | -33.82% | -7.458% | 16.59%
| Z-Score (Absolute) |  | 0.83 | 2 | 6.63 | 0.33 | 0 | 0.23 | 0.18 | 2.14 | 0.48 | 0.67 | 0.07 | 2.18 | 0.19 | 0.76 | 1.52 | 0.47 | 3.09 | 0.65 | 1.58
| Analytical results for Ti in OREAS 611 (Certified Value 0.192 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.1918 | 0.187 | 0.197 | 0.188 | 0.191 | 0.2 | 0.19 | 0.195 | 0.172 | 0.1969 | 0.21055 | 0.1912 | 0.21 | 0.21 | 0.19025 | 0.18 | 0.174 | 0.18 | 0.2 | 0.18 | 0.194 | 0.19 | 0.192 | 1
| 1 | 2 | 0.185 | 0.199 | 0.194 | 0.191 | 0.191 | 0.192 | 0.188 | 0.195 | 0.179 | 0.2022 | 0.20773 | 0.1891 | 0.2 | 0.21 | 0.18885 | 0.18 | 0.182 | 0.17 | 0.2 | 0.17 | 0.195 | 0.2 | 0.1904 | 11
| 1 | 3 | 0.1881 | 0.195 | 0.198 | 0.194 | 0.191 | 0.191 | 0.191 | 0.193 | 0.175 | 0.2041 | 0.20801 | 0.19 | 0.2 | 0.21 | 0.18905 | 0.18 | 0.178 | 0.16 | 0.19 | 0.18 | 0.195 | 0.19 | 0.1988 | 16
| 1 | 4 | 0.1868 | 0.192 | 0.205 | 0.193 | 0.192 | 0.192 | 0.19 | 0.195 | 0.18 | 0.1963 | 0.21038 | 0.19 | 0.2 | 0.21 | 0.18803 | 0.18 | 0.172 | 0.17 | 0.2 | 0.18 | 0.196 | 0.2 | 0.1872 | 0.19187
| 1 | 5 | 0.1912 | 0.194 | 0.203 | 0.198 | 0.191 | 0.183 | 0.187 | 0.202 | 0.173 | 0.198 | 0.20458 | 0.1918 | 0.19 | 0.21 | 0.18879 | 0.18 | 0.18 | 0.17 | 0.19 | 0.16 | 0.194 | 0.2 | 0.1956 | 59
| 1 | 6 | 0.1921 | 0.181 | 0.201 | 0.191 | 0.19 | 0.187 | 0.184 | 0.199 | 0.175 | 0.1966 | 0.20541 | 0.2017 | 0.2 | 0.21 | 0.19125 | 0.18 | 0.178 | 0.16 | 0.2 | 0.18 | 0.194 | 0.19 | 0.194
| Mean |  | 0.18917 | 0.19133 | 0.19967 | 0.1925 | 0.191 | 0.19083 | 0.18833 | 0.1965 | 0.17567 | 0.19902 | 0.20778 | 0.1923 | 0.2 | 0.21 | 0.18937 | 0.18 | 0.17733 | 0.16833 | 0.19667 | 0.175 | 0.19467 | 0.195 | 0.193
| Median |  | 0.18965 | 0.193 | 0.1995 | 0.192 | 0.191 | 0.1915 | 0.189 | 0.195 | 0.175 | 0.19745 | 0.20787 | 0.1906 | 0.2 | 0.21 | 0.18895 | 0.18 | 0.178 | 0.17 | 0.2 | 0.18 | 0.1945 | 0.195 | 0.193
| Std Dev. |  | 0.0029588 | 0.0064083 | 0.0040825 | 0.0033912 | 0.00063246 | 0.0057067 | 0.002582 | 0.0033317 | 0.0032042 | 0.0033078 | 0.0024646 | 0.004704 | 0.0063246 | 0 | 0.0011655 | 3.0405e-17 | 0.0037238 | 0.0075277 | 0.005164 | 0.0083666 | 0.0008165 | 0.0054772 | 0.0040694
| Rel.Std.Dev. |  | 1.564% | 3.349% | 2.045% | 1.762% | 0.3311% | 2.99% | 1.371% | 1.696% | 1.824% | 1.662% | 1.186% | 2.446% | 3.162% | 0% | 0.6155% | 1.689e-14% | 2.1% | 4.472% | 2.626% | 4.781% | 0.4194% | 2.809% | 2.108%
| PDM3 |  | -1.41% | -0.281% | 4.062% | 0.327% | -0.4547% | -0.5416% | -1.845% | 2.412% | -8.446% | 3.723% | 8.289% | 0.2228% | 4.236% | 9.448% | -1.305% | -6.188% | -7.578% | -12.27% | 2.499% | -8.794% | 1.456% | 1.63% | 0.5876%
| Z-Score (Absolute) |  | 0.5 | 0.16 | 1.18 | 0.03 | 0.21 | 0.24 | 0.64 | 0.67 | 2.67 | 1.08 | 2.48 | 0 | 1.24 | 2.84 | 0.47 | 1.97 | 2.4 | 3.85 | 0.7 | 2.78 | 0.38 | 0.43 | 0.11
| Analytical results for Tl in OREAS 611 (Certified Value 2.17 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.1 | 2.09 | 2.31 | 2.12 | 2.16 | 2.18 | 2.12 | 2.09 | 2.18 | 2.22 | 2.3347 | 2.34 | 1.91 | 2.2 | < 5 | 1.7 | 2.13 | 1.98 | 2.3 |  |  |  |  | 1
| 1 | 2 | 2.09 | 2.23 | 2.36 | 2.11 | 2.18 | 2.14 | 2.06 | 2.15 | 2.19 | 2.26 | 2.3015 | 2.07 | 1.96 | 2.1 | < 5 | 1.7 | 2.12 | 1.95 | 2.4 |  |  |  |  | 37
| 1 | 3 | 2.1 | 2.34 | 2.32 | 2.16 | 2.21 | 2.19 | 2.18 | 2.1 | 2.14 | 2.23 | 2.3443 | 2.3 | 2.03 | 2.2 | < 5 | 1.8 | 2.07 | 1.97 | 2.4 |  |  |  |  | 16
| 1 | 4 | 2.14 | 2.19 | 2.36 | 2.15 | 2.15 | 2.13 | 2.22 | 2.07 | 2.18 | 2.16 | 2.2985 | 2.27 | 2.04 | 2.2 | < 5 | 1.7 | 2.1 | 1.97 | 2.4 |  |  |  |  | 2.1746
| 1 | 5 | 2.06 | 2.27 | 2.33 | 2.35 | 2.16 | 2.17 | 2.15 | 2.2 | 2.19 | 2.27 | 2.3066 | 2.17 | 2.02 | 2.2 | < 5 | 1.8 | 2.05 | 1.94 | 2.4 |  |  |  |  | 60
| 1 | 6 | 2.1 | 2.07 | 2.38 | 2.2 | 2.26 | 2.13 | 2.09 | 2.07 | 2.13 | 2.2 | 2.2891 | 2.38 | 2.04 | 2.2 | < 5 | 1.7 | 2.13 | 1.95 | 2.3
| Mean |  | 2.0983 | 2.1983 | 2.3433 | 2.1817 | 2.1867 | 2.1567 | 2.1367 | 2.1133 | 2.1683 | 2.2233 | 2.3124 | 2.255 | 2 | 2.1833 |  | 1.7333 | 2.1 | 1.96 | 2.3667
| Median |  | 2.1 | 2.21 | 2.345 | 2.155 | 2.17 | 2.155 | 2.135 | 2.095 | 2.18 | 2.225 | 2.3041 | 2.285 | 2.025 | 2.2 |  | 1.7 | 2.11 | 1.96 | 2.4
| Std Dev. |  | 0.025626 | 0.10439 | 0.027325 | 0.088412 | 0.041793 | 0.026583 | 0.058878 | 0.05164 | 0.026394 | 0.040332 | 0.021931 | 0.11537 | 0.053292 | 0.040825 |  | 0.05164 | 0.033466 | 0.015492 | 0.05164
| Rel.Std.Dev. |  | 1.221% | 4.748% | 1.166% | 4.052% | 1.911% | 1.233% | 2.756% | 2.444% | 1.217% | 1.814% | 0.9484% | 5.116% | 2.665% | 1.87% |  | 2.979% | 1.594% | 0.7904% | 2.182%
| PDM3 |  | -3.508% | 1.091% | 7.759% | 0.3243% | 0.5542% | -0.8253% | -1.745% | -2.818% | -0.2888% | 2.24% | 6.338% | 3.697% | -8.03% | 0.4009% |  | -20.29% | -3.431% | -9.869% | 8.832%
| Z-Score (Absolute) |  | 0.77 | 0.15 | 1.49 | 0 | 0.05 | 0.23 | 0.41 | 0.63 | 0.12 | 0.38 | 1.2 | 0.67 | 1.67 | 0.02 | 2.93 | 4.12 | 0.75 | 2.04 | 1.7
| Analytical results for Tm in OREAS 611 (Indicative Value 91.7 ppb) | [24 empty] | Indicative
Tm | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab |  |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16 |  |  |  |  | ppb
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100 |  |  |  |  | 1
| 1 | 2 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100 |  |  |  |  | 6
| 1 | 3 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 80 |  |  |  |  | 16
| 1 | 4 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100 |  |  |  |  | 91.6667
| 1 | 5 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 100 |  |  |  |  | 12
| 1 | 6 | 2.02: < 100; 2.03: < 100; 2.04: 90; 2.05: < 100; 2.06: 80
| Mean |  | 2.04: 90; 2.06: 93.3333
| Median |  | 2.04: 90; 2.06: 100
| Std Dev. |  | 2.04: 0; 2.06: 10.328
| Rel.Std.Dev. |  | 2.04: 0%; 2.06: 11.07%
| PDM3 |  | 2.04: -1.818%; 2.06: 1.818%
Z-Score (Absolute) | bar: 2.02 to 2.06
| Analytical results for U in OREAS 611 (Certified Value 2.68 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.56 | 2.7 | 2.7 | 2.7 | 2.5 | 2.6 | 2.5 | 2.5 | 2.8 | 2.78 | 2.8165 | 3 | 2.57 | 2.6 | < 10 | 1.96 | 2.5 | 2.48 | 3 |  |  |  |  | 1
| 1 | 2 | 2.5 | 2.7 | 2.8 | 2.8 | 2.5 | 2.6 | 2.4 | 2.6 | 2.89 | 2.72 | 2.7606 | 2.7 | 2.63 | 2.6 | < 10 | 2.04 | 2.6 | 2.49 | 3.1 |  |  |  |  | 39
| 1 | 3 | 2.56 | 2.6 | 2.8 | 2.7 | 2.5 | 2.7 | 2.6 | 2.5 | 2.71 | 2.77 | 2.8398 | 2.9 | 2.63 | 2.6 | < 10 | 2.09 | 2.5 | 2.48 | 3.1 |  |  |  |  | 16
| 1 | 4 | 2.61 | 2.9 | 2.8 | 2.7 | 2.5 | 2.6 | 2.7 | 2.5 | 2.88 | 2.75 | 2.7271 | 2.9 | 2.7 | 2.5 | < 10 | 1.99 | 2.4 | 2.49 | 3 |  |  |  |  | 2.6771
| 1 | 5 | 2.6 | 2.7 | 2.9 | 2.9 | 2.5 | 2.5 | 2.6 | 2.6 | 2.87 | 2.81 | 2.7671 | 2.6 | 2.69 | 2.6 | < 10 | 2.02 | 2.5 | 2.38 | 3 |  |  |  |  | 61
| 1 | 6 | 2.67 | 2.5 | 2.9 | 2.8 | 2.7 | 2.5 | 2.6 | 2.5 | 2.85 | 2.78 | 2.7549 | 2.9 | 2.59 | 2.6 | < 10 | 2.02 | 2.4 | 2.5 | 2.9
| Mean |  | 2.5833 | 2.6833 | 2.8167 | 2.7667 | 2.5333 | 2.5833 | 2.5667 | 2.5333 | 2.8333 | 2.7683 | 2.7777 | 2.8333 | 2.635 | 2.5833 |  | 2.02 | 2.4833 | 2.47 | 3.0167
| Median |  | 2.58 | 2.7 | 2.8 | 2.75 | 2.5 | 2.6 | 2.6 | 2.5 | 2.86 | 2.775 | 2.7638 | 2.9 | 2.63 | 2.6 |  | 2.02 | 2.5 | 2.485 | 3
| Std Dev. |  | 0.057504 | 0.13292 | 0.075277 | 0.08165 | 0.08165 | 0.075277 | 0.10328 | 0.05164 | 0.068313 | 0.030605 | 0.042077 | 0.15055 | 0.052058 | 0.040825 |  | 0.044272 | 0.075277 | 0.044721 | 0.075277
| Rel.Std.Dev. |  | 2.226% | 4.953% | 2.673% | 2.951% | 3.223% | 2.914% | 4.024% | 2.038% | 2.411% | 1.106% | 1.515% | 5.314% | 1.976% | 1.58% |  | 2.192% | 3.031% | 1.811% | 2.495%
| PDM3 |  | -3.502% | 0.2337% | 5.214% | 3.347% | -5.369% | -3.502% | -4.124% | -5.369% | 5.837% | 3.409% | 3.757% | 5.837% | -1.572% | -3.502% |  | -24.54% | -7.237% | -7.735% | 12.69%
| Z-Score (Absolute) |  | 0.26 | 0.24 | 0.92 | 0.67 | 0.51 | 0.26 | 0.35 | 0.51 | 1 | 0.67 | 0.72 | 1 | 0 | 0.26 | 11.96 | 3.11 | 0.77 | 0.83 | 1.93
| Analytical results for V in OREAS 611 (Certified Value 31.9 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 33 | 32 | 28 | 31 | 32 | 33 | 31 | 33 | 31 | 33 | 31.7472 | 33 | 34.7 | 35 | 32.3962 | 29 | 31.751 | 34 | 33 | 29 | 35 | 33 | 38.16 | 1
| 1 | 2 | 32 | 33 | 29 | 32 | 31 | 32 | 32 | 33 | 32 | 34 | 31.5189 | 29 | 33.4 | 40 | 32.5193 | 29 | 32.339 | 33 | 33 | 29 | 36 | 33 | 38.22 | 40
| 1 | 3 | 32 | 33 | 29 | 33 | 31 | 32 | 31 | 33 | 31 | 35 | 32.1475 | 32 | 34 | 40 | 32.6576 | 29 | 32.042 | 32 | 33 | 29 | 36 | 33 | 38.48 | 16
| 1 | 4 | 32 | 33 | 29 | 32 | 31 | 32 | 31 | 33 | 32 | 34 | 31.7392 | 29 | 33.5 | 35 | 31.9668 | 29 | 30.929 | 33 | 33 | 30 | 36 | 33 | 38.05 | 31.8691
| 1 | 5 | 31 | 33 | 31 | 33 | 32 | 31 | 31 | 34 | 31 | 34 | 31.8456 | 30 | 33.3 | 40 | 32.4542 | 30 | 32.389 | 32 | 32 | 29 | 35 | 33 | 38.8 | 62
| 1 | 6 | 33 | 31 | 31 | 32 | 32 | 32 | 31 | 33 | 31 | 33 | 32.0943 | 34 | 33.9 | 40 | 32.3748 | 29 | 31.581 | 33 | 33 | 29 | 34 | 33 | 38.46
| Mean |  | 32.1667 | 32.5 | 29.5 | 32.1667 | 31.5 | 32 | 31.1667 | 33.1667 | 31.3333 | 33.8333 | 31.8488 | 31.1667 | 33.8 | 38.3333 | 32.3948 | 29.1667 | 31.8385 | 32.8333 | 32.8333 | 29.1667 | 35.3333 | 33 | 38.3617
| Median |  | 32 | 33 | 29 | 32 | 31.5 | 32 | 31 | 33 | 31 | 34 | 31.7964 | 31 | 33.7 | 40 | 32.4252 | 29 | 31.8965 | 33 | 33 | 29 | 35.5 | 33 | 38.34
| Std Dev. |  | 0.75277 | 0.83666 | 1.2247 | 0.75277 | 0.54772 | 0.63246 | 0.40825 | 0.40825 | 0.5164 | 0.75277 | 0.2369 | 2.137 | 0.52154 | 2.582 | 0.23309 | 0.40825 | 0.54701 | 0.75277 | 0.40825 | 0.40825 | 0.8165 | 0 | 0.27338
| Rel.Std.Dev. |  | 2.34% | 2.574% | 4.152% | 2.34% | 1.739% | 1.976% | 1.31% | 1.231% | 1.648% | 2.225% | 0.7438% | 6.857% | 1.543% | 6.736% | 0.7195% | 1.4% | 1.718% | 2.293% | 1.243% | 1.4% | 2.311% | 0% | 0.7126%
| PDM3 |  | 0.9337% | 1.98% | -7.434% | 0.9337% | -1.158% | 0.4107% | -2.204% | 4.072% | -1.681% | 6.163% | -0.06374% | -2.204% | 6.059% | 20.28% | 1.65% | -8.48% | -0.09603% | 3.026% | 3.026% | -8.48% | 10.87% | 3.549% | 20.37%
| Z-Score (Absolute) |  | 0 | 0.27 | 2.16 | 0 | 0.54 | 0.13 | 0.81 | 0.81 | 0.67 | 1.35 | 0.26 | 0.81 | 1.32 | 4.99 | 0.18 | 2.43 | 0.27 | 0.54 | 0.54 | 2.43 | 2.56 | 0.67 | 5.01
| Analytical results for W in OREAS 611 (Certified Value 8.75 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 8.7 | 8.9 | 6.7 | 9.2 | 8.9 | 8.9 | 8.4 | 9 | 8.5 | 8.8 | 8.2897 | 9.3 | 10.3 | 8.5 | < 20 | 7.7 | < 10 | 8.6 | 8.6 | 17 | 10 | < 5 |  | 1
| 1 | 2 | 8.4 | 8.8 | 6.6 | 8.7 | 9.1 | 8.5 | 8.3 | 9.1 | 8.9 | 9 | 8.9134 | 8.4 | 10 | 8.5 | < 20 | 8.2 | < 10 | 8.6 | 8.4 | 17 | 9.7 | < 5 |  | 41
| 1 | 3 | 8.6 | 9 | 6.8 | 9.1 | 8.7 | 8.6 | 8.6 | 8.8 | 8.4 | 9 | 8.7326 | 9.6 | 9.9 | 8.5 | < 20 | 8.1 | < 10 | 8.5 | 8.3 | 15 | 9.9 | < 5 |  | 16
| 1 | 4 | 8.7 | 9.1 | 6.9 | 8.9 | 8.6 | 8.6 | 9.1 | 8.8 | 8.7 | 8.9 | 8.5224 | 9.7 | 10.2 | 8.5 | < 20 | 8 | < 10 | 8.2 | 8.5 | 16 | 9.6 | < 5 |  | 8.7494
| 1 | 5 | 8.7 | 8.6 | 6.2 | 9.3 | 8.7 | 8.7 | 8.6 | 9.2 | 8.7 | 9.1 | 8.1375 | 8.4 | 10.1 | 8.5 | < 20 | 8.4 | < 10 | 8.4 | 8 | 15 | 9.5 | < 5 |  | 63
| 1 | 6 | 8.8 | 8.1 | 6.3 | 9.3 | 8.9 | 8.6 | 8.8 | 9 | 8.7 | 9.2 | 8.3423 | 9.9 | 9.9 | 8.5 | < 20 | 7.9 | < 10 | 8.3 | 8.4 | 15 | 9.6 | < 5
| Mean |  | 8.65 | 8.75 | 6.5833 | 9.0833 | 8.8167 | 8.65 | 8.6333 | 8.9833 | 8.65 | 9 | 8.4897 | 9.2167 | 10.0667 | 8.5 |  | 8.05 |  | 8.4333 | 8.3667 | 15.8333 | 9.7167
| Median |  | 8.7 | 8.85 | 6.65 | 9.15 | 8.8 | 8.6 | 8.6 | 9 | 8.7 | 9 | 8.4324 | 9.45 | 10.05 | 8.5 |  | 8.05 |  | 8.45 | 8.4 | 15.5 | 9.65
| Std Dev. |  | 0.13784 | 0.36194 | 0.27869 | 0.24014 | 0.18348 | 0.13784 | 0.28752 | 0.16021 | 0.17607 | 0.14142 | 0.29169 | 0.66156 | 0.1633 | 0 |  | 0.2429 |  | 0.1633 | 0.20656 | 0.98319 | 0.19408
| Rel.Std.Dev. |  | 1.594% | 4.136% | 4.233% | 2.644% | 2.081% | 1.594% | 3.33% | 1.783% | 2.035% | 1.571% | 3.436% | 7.178% | 1.622% | 0% |  | 3.017% |  | 1.936% | 2.469% | 6.21% | 1.997%
| PDM3 |  | -1.136% | 0.007383% | -24.76% | 3.817% | 0.7693% | -1.136% | -1.326% | 2.674% | -1.136% | 2.865% | -2.968% | 5.341% | 15.06% | -2.85% |  | -7.993% |  | -3.612% | -4.374% | 80.97% | 11.06%
| Z-Score (Absolute) |  | 0 | 0.2 | 4.08 | 0.86 | 0.33 | 0 | 0.03 | 0.66 | 0 | 0.69 | 0.32 | 1.12 | 2.8 | 0.3 | 2.66 | 1.18 | 7.2 | 0.43 | 0.56 | 14.18 | 2.11 | 12.14
| Analytical results for Y in OREAS 611 (Certified Value 6.8 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.9 | 7.4 | 6.7 | 6.5 | 6.6 | 7 | 7 | 7.4 | 6.9 | 6.82 | 6.8339 | 6.8 | 7 | 6.8 | 7.0472 | 8.8 | 6.367 | 5.4 | 6.5 | 6 | 7.2 |  |  | 1
| 1 | 2 | 6.9 | 7.1 | 6.8 | 6.7 | 6.3 | 6.9 | 6.7 | 7 | 6.6 | 6.7 | 6.6805 | 6.4 | 7 | 6.9 | 7.3447 | 8.9 | 6.324 | 5.8 | 6.4 | 5.9 | 7.3 |  |  | 42
| 1 | 3 | 7.1 | 6.6 | 6.6 | 6.4 | 6.3 | 6.8 | 6.8 | 7.3 | 6.7 | 6.78 | 6.9018 | 6.5 | 6.9 | 6.8 | 7.2195 | 9.4 | 6.602 | 5.6 | 6.5 | 6 | 7.1 |  |  | 16
| 1 | 4 | 7 | 7.2 | 6.8 | 6.5 | 6.4 | 6.9 | 6.9 | 7.4 | 7 | 6.7 | 6.4478 | 5.8 | 7.1 | 7.1 | 7.0087 | 9.2 | 6.583 | 5.6 | 6.5 | 6.1 | 6.9 |  |  | 6.8039
| 1 | 5 | 7.2 | 7.1 | 7.1 | 7.3 | 6.5 | 6.7 | 6.8 | 7.8 | 6.7 | 6.73 | 6.6171 | 6 | 6.9 | 6.9 | 6.9705 | 9.3 | 5.92 | 5.4 | 6.3 | 6 | 6.7 |  |  | 64
| 1 | 6 | 7.3 | 6.6 | 6.9 | 6.8 | 6.6 | 6.7 | 6.7 | 7.6 | 6.9 | 6.77 | 6.6508 | 6.7 | 7 | 6.8 | 7.1592 | 9.2 | 6.045 | 5.7 | 6.5 | 5.8 | 6.9
| Mean |  | 7.0667 | 7 | 6.8167 | 6.7 | 6.45 | 6.8333 | 6.8167 | 7.4167 | 6.8 | 6.75 | 6.6886 | 6.3667 | 6.9833 | 6.8833 | 7.125 | 9.1333 | 6.3068 | 5.5833 | 6.45 | 5.9667 | 7.0167
| Median |  | 7.05 | 7.1 | 6.8 | 6.6 | 6.45 | 6.85 | 6.8 | 7.4 | 6.8 | 6.75 | 6.6657 | 6.45 | 7 | 6.85 | 7.1032 | 9.2 | 6.3455 | 5.6 | 6.5 | 6 | 7
| Std Dev. |  | 0.1633 | 0.32863 | 0.17224 | 0.32863 | 0.13784 | 0.12111 | 0.1169 | 0.27142 | 0.15492 | 0.048166 | 0.16196 | 0.39328 | 0.075277 | 0.1169 | 0.1427 | 0.23381 | 0.27767 | 0.16021 | 0.083666 | 0.10328 | 0.22286
| Rel.Std.Dev. |  | 2.311% | 4.695% | 2.527% | 4.905% | 2.137% | 1.772% | 1.715% | 3.66% | 2.278% | 0.7136% | 2.421% | 6.177% | 1.078% | 1.698% | 2.003% | 2.56% | 4.403% | 2.869% | 1.297% | 1.731% | 3.176%
| PDM3 |  | 3.862% | 2.882% | 0.1874% | -1.527% | -5.202% | 0.4324% | 0.1874% | 9.006% | -0.05751% | -0.7924% | -1.694% | -6.426% | 2.637% | 1.167% | 4.719% | 34.24% | -7.306% | -17.94% | -5.202% | -12.31% | 3.127%
| Z-Score (Absolute) |  | 0.84 | 0.62 | 0 | 0.39 | 1.24 | 0.06 | 0 | 2.02 | 0.06 | 0.22 | 0.43 | 1.52 | 0.56 | 0.22 | 1.04 | 7.81 | 1.72 | 4.16 | 1.24 | 2.87 | 0.67
| Analytical results for Yb in OREAS 611 (Certified Value 0.55 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.11; 2.05: 2.14; 2.06: 2.16; 2.07: 2.18; 2.08: 2.20; 2.09: 2.21 |  |  | ppm
|  |  | 2.02: 4A*OES/MS; 2.03: 4A*OES/MS; 2.04: 4A*MS; 2.05: 4A*MS; 2.06: 4A*MS; 2.07: 4A*OES/MS; 2.08: 4A*OES/MS; 2.09: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 2.02: 0.5; 2.03: 0.6; 2.04: 0.59; 2.05: 0.5; 2.06: 0.65; 2.07: 0.4; 2.08: 0.5; 2.09: 0.5 |  |  | 1
| 1 | 2 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.58; 2.05: 0.5; 2.06: 0.7; 2.07: 0.4; 2.08: 0.4; 2.09: 0.5 |  |  | 13
| 1 | 3 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.61; 2.05: 0.6; 2.06: 0.7; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5 |  |  | 16
| 1 | 4 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.59; 2.05: 0.6; 2.06: 0.7; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5 |  |  | 0.55042
| 1 | 5 | 2.02: 0.6; 2.03: 0.6; 2.04: 0.61; 2.05: 0.5; 2.06: 0.65; 2.07: 0.5; 2.08: 0.4; 2.09: 0.5 |  |  | 65
| 1 | 6 | 2.02: 0.5; 2.03: 0.6; 2.04: 0.59; 2.05: 0.5; 2.06: 0.65; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5
| Mean |  | 2.02: 0.56667; 2.03: 0.6; 2.04: 0.595; 2.05: 0.53333; 2.06: 0.675; 2.07: 0.46667; 2.08: 0.46667; 2.09: 0.5
| Median |  | 2.02: 0.6; 2.03: 0.6; 2.04: 0.59; 2.05: 0.5; 2.06: 0.675; 2.07: 0.5; 2.08: 0.5; 2.09: 0.5
| Std Dev. |  | 2.02: 0.05164; 2.03: 0; 2.04: 0.012247; 2.05: 0.05164; 2.06: 0.027386; 2.07: 0.05164; 2.08: 0.05164; 2.09: 0
| Rel.Std.Dev. |  | 2.02: 9.113%; 2.03: 0%; 2.04: 2.058%; 2.05: 9.682%; 2.06: 4.057%; 2.07: 11.07%; 2.08: 11.07%; 2.09: 0%
| PDM3 |  | 2.02: 2.952%; 2.03: 9.008%; 2.04: 8.1%; 2.05: -3.104%; 2.06: 22.63%; 2.07: -15.22%; 2.08: -15.22%; 2.09: -9.16%
| Z-Score (Absolute) |  | 2.02: 0.22; 2.03: 0.67; 2.04: 0.61; 2.05: 0.22; 2.06: 1.69; 2.07: 1.12; 2.08: 1.12; 2.09: 0.67
| Analytical results for Zn in OREAS 611 (Certified Value 2023 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 2077 | 1936.9 | 1870 | 1960 | 2030 | 2030 | 2010 | 2070 | 2090 | 1933 | 2050.6204 | 2191 | 2030 | 1960 | 2051.5522 | 1975 | 2106.495 | 1935 | 2254 | 1947 | 2148 | 1970 | 1975 | 1
| 1 | 2 | 2025 | 1935.5 | 1960 | 1970 | 2010 | 1990 | 1990 | 2120 | 2200 | 1955 | 2067.8706 | 2062.1 | 2070 | 1980 | 2072.5873 | 1996 | 2102.149 | 1936 | 2246 | 1864 | 2152 | 2000 | 1994 | 22
| 1 | 3 | 2059 | 1963.9 | 1950 | 2020 | 2020 | 1980 | 2040 | 2090 | 2070 | 1975 | 2086.664 | 2078.6 | 2040 | 2010 | 2039.6767 | 2020 | 2123.744 | 1897 | 2162 | 1962 | 2159 | 1970 | 1984 | 16
| 1 | 4 | 1959 | 1957.5 | 2000 | 1980 | 2030 | 1980 | 2050 | 2100 | 2160 | 1946 | 2030.4674 | 2072.8 | 2090 | 2000 | 2053.4513 | 1981 | 2084.687 | 1885 | 2209 | 1971 | 2163 | 2020 | 1998 | 2022.9122
| 1 | 5 | 2047 | 1987.4 | 2070 | 2040 | 2010 | 1900 | 2010 | 2180 | 2130 | 1958 | 2075.3935 | 2074 | 2040 | 1960 | 2067.431 | 1977 | 2097.638 | 1887 | 2252 | 1871 | 2160 | 2010 | 2001 | 66
| 1 | 6 | 2062 | 1841 | 2060 | 1980 | 1960 | 1940 | 1990 | 2150 | 2150 | 1957 | 2010.8731 | 2305 | 2040 | 2010 | 2054.2421 | 1983 | 2070.233 | 1888 | 2234 | 2054 | 2159 | 1960 | 1997
| Mean |  | 2038.1667 | 1937.0333 | 1985 | 1991.6667 | 2010 | 1970 | 2015 | 2118.3333 | 2133.3333 | 1954 | 2053.6482 | 2130.5833 | 2051.6667 | 1986.6667 | 2056.4901 | 1988.6667 | 2097.491 | 1904.6667 | 2226.1667 | 1944.8333 | 2156.8333 | 1988.3333 | 1991.5
| Median |  | 2053 | 1947.2 | 1980 | 1980 | 2015 | 1980 | 2010 | 2110 | 2140 | 1956 | 2059.2455 | 2076.3 | 2040 | 1990 | 2053.8467 | 1982 | 2099.8935 | 1892.5 | 2240 | 1954.5 | 2159 | 1985 | 1995.5
| Std Dev. |  | 42.4943 | 50.7982 | 75.0333 | 31.2517 | 26.0768 | 44.7214 | 25.0998 | 40.7022 | 47.6095 | 13.9427 | 28.7798 | 97.9838 | 23.1661 | 23.3809 | 11.8337 | 17.0255 | 18.4334 | 24.2377 | 35.5214 | 70.5505 | 5.6362 | 24.8328 | 9.975
| Rel.Std.Dev. |  | 2.085% | 2.622% | 3.78% | 1.569% | 1.297% | 2.27% | 1.246% | 1.921% | 2.232% | 0.7135% | 1.401% | 4.599% | 1.129% | 1.177% | 0.5754% | 0.8561% | 0.8788% | 1.273% | 1.596% | 3.628% | 0.2613% | 1.249% | 0.5009%
| PDM3 |  | 0.7541% | -4.245% | -1.874% | -1.545% | -0.6383% | -2.616% | -0.3911% | 4.717% | 5.459% | -3.407% | 1.519% | 5.323% | 1.421% | -1.792% | 1.66% | -1.693% | 3.687% | -5.845% | 10.05% | -3.86% | 6.62% | -1.709% | -1.553%
| Z-Score (Absolute) |  | 0.44 | 1.13 | 0.39 | 0.28 | 0 | 0.62 | 0.08 | 1.67 | 1.91 | 0.87 | 0.67 | 1.86 | 0.64 | 0.36 | 0.72 | 0.33 | 1.35 | 1.63 | 3.34 | 1.01 | 2.27 | 0.33 | 0.29
| Analytical results for Zr in OREAS 611 (Certified Value 69 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.16 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 65 | 65.1 | 82 | 67 | 71.8 | 71.6 | 71.9 | 74.4 | 71.1 | 72.1 | 72.6632 | 64 | 75.1 | 74 | 79.6701 | 61 | 59.755 | 72.4 | 64.6 | 61.4 | 68.5 | 61 | 87.17 | 1
| 1 | 2 | 62.9 | 66.1 | 88 | 70.2 | 71.1 | 69.8 | 68.3 | 72.9 | 71.8 | 75.7 | 72.2543 | 72.9 | 72.3 | 75 | 77.9222 | 63 | 61.365 | 72.1 | 64.8 | 59.7 | 72.3 | 62 | 91.51 | 23
| 1 | 3 | 66.1 | 66 | 80 | 68 | 71.4 | 70.2 | 70.9 | 72.8 | 73.2 | 74.5 | 73.5216 | 75.8 | 73.7 | 77 | 76.2772 | 61 | 59.586 | 68.5 | 63.6 | 59.5 | 69.9 | 61 | 88.85 | 16
| 1 | 4 | 64.9 | 70.2 | 85 | 68.3 | 71.9 | 69.5 | 72.5 | 72.1 | 71.5 | 72.5 | 70.7372 | 73.2 | 72.4 | 76 | 73.957 | 59 | 57.312 | 67.6 | 66.9 | 62 | 68.1 | 61 | 86.19 | 68.8064
| 1 | 5 | 66.5 | 67 | 84 | 70.3 | 72.3 | 67.6 | 70.3 | 75.9 | 72.1 | 71.9 | 70.8256 | 74 | 72.1 | 73 | 77.921 | 64 | 58.137 | 67.9 | 64.6 | 59.8 | 68.4 | 61 | 87.53 | 67
| 1 | 6 | 66.1 | 62.8 | 89 | 72.7 | 68 | 67.4 | 68.9 | 75.8 | 72.3 | 69.7 | 71.3178 | 67 | 72.2 | 75 | 78.372 | 62 | 58.299 | 70.3 | 65.4 | 59 | 65.9 | 60 | 89.87
| Mean |  | 65.25 | 66.2 | 84.6667 | 69.4167 | 71.0833 | 69.35 | 70.4667 | 73.9833 | 72 | 72.7333 | 71.8866 | 71.15 | 72.9667 | 75 | 77.3532 | 61.6667 | 59.0757 | 69.8 | 64.9833 | 60.2333 | 68.85 | 61 | 88.52
| Median |  | 65.55 | 66.05 | 84.5 | 69.25 | 71.6 | 69.65 | 70.6 | 73.65 | 71.95 | 72.3 | 71.786 | 73.05 | 72.35 | 75 | 77.9216 | 61.5 | 58.9425 | 69.4 | 64.7 | 59.75 | 68.45 | 61 | 88.19
| Std Dev. |  | 1.3202 | 2.4273 | 3.4448 | 2.0624 | 1.5664 | 1.6047 | 1.6464 | 1.629 | 0.72664 | 2.1106 | 1.112 | 4.5903 | 1.1994 | 1.4142 | 1.9872 | 1.7512 | 1.4531 | 2.1185 | 1.1035 | 1.1843 | 2.1258 | 0.63246 | 1.9546
| Rel.Std.Dev. |  | 2.023% | 3.667% | 4.069% | 2.971% | 2.204% | 2.314% | 2.336% | 2.202% | 1.009% | 2.902% | 1.547% | 6.452% | 1.644% | 1.886% | 2.569% | 2.84% | 2.46% | 3.035% | 1.698% | 1.966% | 3.088% | 1.037% | 2.208%
| PDM3 |  | -5.169% | -3.788% | 23.05% | 0.8869% | 3.309% | 0.79% | 2.413% | 7.524% | 4.641% | 5.707% | 4.477% | 3.406% | 6.046% | 9.001% | 12.42% | -10.38% | -14.14% | 1.444% | -5.556% | -12.46% | 0.06334% | -11.35% | 28.65%
| Z-Score (Absolute) |  | 1 | 0.82 | 2.72 | 0.2 | 0.12 | 0.21 | 0 | 0.67 | 0.29 | 0.43 | 0.27 | 0.13 | 0.48 | 0.87 | 1.32 | 1.69 | 2.18 | 0.13 | 1.05 | 1.96 | 0.31 | 1.82 | 3.46
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C562: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C892: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C964: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1036: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 611 (Certified Value 79.2 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 1
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 2
| 1 | 1 | 80.1 | 81.1 | 81.026 | 72.2 | 76.4 | 82.2 | 79 | 78.9 | 79 | 80.7 | 77.49 | 80.6336 | 57.6 | 75.876 | 73.2 | 81.052 | 79.7 | 79 | > 10 | 81.37 | 74 | 86.1 | 73 | 79.3 | 1
| 1 | 2 | 79.3 | 86.3 | 80.612 | 75.7 | 76.7 | 82.1 | 78.9 | 78.9 | 80.3 | 80.3 | 76.04 | 80.4652 | 57.4 | 75.399 | 72.2 | 80.7914 | 80 | 82 | > 10 | 81.61 | 72 | 86 | 74 | 79.64 | 29
| 1 | 3 | 79.3 | 87.4 | 81.143 | 75.3 | 77.3 | 83.1 | 79.3 | 78.9 | 80.4 | 79.7 | 76.09 | 80.5958 | 57.5 | 77.769 | 72.3 | 81.4209 | 79.9 | 81 | > 10 | 81.04 | 76 | 85.5 | 73 | 79.98 | 16
| 1 | 4 | 79.5 | 86.2 | 82.034 | 72.8 | 77.3 | 83.5 | 78.9 | 79.5 | 79.9 | 80 | 77.09 | 80.1227 | 57.5 | 79.024 | 74.4 | 81.1491 | 80.1 | 82 | > 10 | 82.05 | 76 | 86.9 | 74 | 80.32 | 79.2496
| 1 | 5 | 82.6 | 85.1 | 79.814 | 72.7 | 77.4 | 84.6 | 80.3 | 79.5 | 80.1 | 79.6 | 75.62 | 80.5854 | 56.5 | 77.642 | 73.6 | 81.3044 | 78.5 | 81 | > 10 | 81.44 | 75 | 89.8 | 74 | 80.66 | 69
| 1 | 6 | 78 | 86.8 | 81.844 | 73.2 | 78.2 | 79.8 | 79.8 | 76.8 | 78.7 | 79.2 | 75.76 | 81.4006 | 55.6 | 77.008 | 73.3 | 79.111 | 79.6 | 82 | > 10 | 83.96 | 74 | 96 | 73 | 80
| Mean |  | 79.8 | 85.4833 | 81.0788 | 73.65 | 77.2167 | 82.55 | 79.3667 | 78.75 | 79.7333 | 79.9167 | 76.3483 | 80.6339 | 57.0167 | 77.1197 | 73.1667 | 80.8048 | 79.6333 | 81.1667 |  | 81.9117 | 74.5 | 88.3833 | 73.5 | 79.9833
| Median |  | 79.4 | 86.25 | 81.0845 | 73 | 77.3 | 82.65 | 79.15 | 78.9 | 80 | 79.85 | 76.065 | 80.5906 | 57.45 | 77.325 | 73.25 | 81.1006 | 79.8 | 81.5 |  | 81.525 | 74.5 | 86.5 | 73.5 | 79.99
| Std Dev. |  | 1.5336 | 2.2781 | 0.81516 | 1.4734 | 0.62423 | 1.6306 | 0.57155 | 0.9995 | 0.71181 | 0.53448 | 0.76051 | 0.41986 | 0.80353 | 1.3297 | 0.82624 | 0.85766 | 0.58538 | 1.169 |  | 1.0565 | 1.5166 | 4.0356 | 0.54772 | 0.4809
| Rel.Std.Dev. |  | 1.922% | 2.665% | 1.005% | 2.001% | 0.8084% | 1.975% | 0.7201% | 1.269% | 0.8927% | 0.6688% | 0.9961% | 0.5207% | 1.409% | 1.724% | 1.129% | 1.061% | 0.7351% | 1.44% |  | 1.29% | 2.036% | 4.566% | 0.7452% | 0.6013%
| PDM3 |  | 0.6946% | 7.866% | 2.308% | -7.066% | -2.565% | 4.165% | 0.1478% | -0.6304% | 0.6104% | 0.8418% | -3.661% | 1.747% | -28.05% | -2.688% | -7.676% | 1.962% | 0.4842% | 2.419% |  | 3.359% | -5.993% | 11.53% | -7.255% | 0.9259%
| Z-Score (Absolute) |  | 0.02 | 1.78 | 0.42 | 1.88 | 0.78 | 0.87 | 0.11 | 0.3 | 0 | 0.06 | 1.05 | 0.28 | 7.03 | 0.81 | 2.03 | 0.33 | 0.03 | 0.44 | NA | 0.67 | 1.62 | 2.68 | 1.93 | 0.08
| Analytical results for Al in OREAS 611 (Certified Value 0.838 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 1.0838 | 0.93 | 1.3 | 0.8 | 0.87 | 0.85 | 0.86 | 0.83 | 1.15 | 0.751 | 1.1336 | 0.61 | 0.82 | 0.65 | 0.83639 | 0.74 | 0.889 | 0.83 | 0.8 | 0.82 | 0.78 | 0.8 | 0.8324 |  | 1
| 1 | 2 | 1.1423 | 0.91 | 1.32 | 0.83 | 0.88 | 0.86 | 0.86 | 0.85 | 1.18 | 0.856 | 1.1428 | 0.61 | 0.8 | 0.64 | 0.82809 | 0.74 | 0.884 | 0.82 | 0.82 | 0.85 | 0.78 | 0.8 | 0.8338 |  | #N/A
| 1 | 3 | 1.2025 | 0.93 | 1.34 | 0.83 | 0.89 | 0.86 | 0.85 | 0.86 | 1.12 | 0.863 | 1.1369 | 0.61 | 0.86 | 0.66 | 0.81986 | 0.73 | 0.888 | 0.83 | 0.86 | 0.85 | 0.78 | 0.79 | 0.8352 |  | 16
| 1 | 4 | 1.1644 | 0.95 | 1.29 | 0.82 | 0.91 | 0.85 | 0.88 | 0.83 | 1.15 | 0.885 | 1.1272 | 0.61 | 0.86 | 0.69 | 0.81395 | 0.75 | 0.883 | 0.84 | 0.83 | 0.83 | 0.79 | 0.8 | 0.8366 |  | 0.83827
| 1 | 5 | 1.1378 | 0.95 | 1.29 | 0.83 | 0.9 | 0.91 | 0.87 | 0.84 | 1.14 | 0.855 | 1.152 | 0.61 | 0.85 | 0.67 | 0.84241 | 0.72 | 0.892 | 0.83 | 0.86 | 0.83 | 0.82 | 0.79 | 0.838 |  | 70
| 1 | 6 | 1.1439 | 0.94 | 1.3 | 0.86 | 0.85 | 0.9 | 0.81 | 0.82 | 1.17 | 0.734 | 1.1442 | 0.59 | 0.84 | 0.64 | 0.7917 | 0.74 | 0.904 | 0.85 | 0.85 | 0.85 | 0.76 | 0.8 | 0.8394
| Mean |  | 1.1458 | 0.935 | 1.3067 | 0.82833 | 0.88333 | 0.87167 | 0.855 | 0.83833 | 1.1517 | 0.824 | 1.1395 | 0.60667 | 0.83833 | 0.65833 | 0.82207 | 0.73667 | 0.89 | 0.83333 | 0.83667 | 0.83833 | 0.785 | 0.79667 | 0.8359
| Median |  | 1.1431 | 0.935 | 1.3 | 0.83 | 0.885 | 0.86 | 0.86 | 0.835 | 1.15 | 0.8555 | 1.1399 | 0.61 | 0.845 | 0.655 | 0.82398 | 0.74 | 0.8885 | 0.83 | 0.84 | 0.84 | 0.78 | 0.8 | 0.8359
| Std Dev. |  | 0.03869 | 0.015166 | 0.019664 | 0.019408 | 0.021602 | 0.026394 | 0.02429 | 0.01472 | 0.02137 | 0.064274 | 0.0087297 | 0.008165 | 0.024014 | 0.019408 | 0.018156 | 0.010328 | 0.0076158 | 0.010328 | 0.024221 | 0.013292 | 0.019748 | 0.005164 | 0.0026192
| Rel.Std.Dev. |  | 3.377% | 1.622% | 1.505% | 2.343% | 2.446% | 3.028% | 2.841% | 1.756% | 1.856% | 7.8% | 0.7661% | 1.346% | 2.864% | 2.948% | 2.209% | 1.402% | 0.8557% | 1.239% | 2.895% | 1.585% | 2.516% | 0.6482% | 0.3133%
| PDM3 |  | 36.68% | 11.54% | 55.88% | -1.186% | 5.375% | 3.984% | 1.995% | 0.007238% | 37.39% | -1.703% | 35.93% | -27.63% | 0.007238% | -21.47% | -1.933% | -12.12% | 6.171% | -0.5892% | -0.1916% | 0.007238% | -6.355% | -4.963% | -0.283%
| Z-Score (Absolute) |  | 4.98 | 1.56 | 7.58 | 0.16 | 0.73 | 0.54 | 0.27 | 0 | 5.07 | 0.23 | 4.87 | 3.75 | 0 | 2.91 | 0.26 | 1.65 | 0.84 | 0.08 | 0.03 | 0 | 0.86 | 0.67 | 0.04
| Analytical results for As in OREAS 611 (Certified Value 3340 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 0
| 1 | 1 | 2580 | 2978 | 3626.1 | 3120 | 3390 | 3640 | 3500 | 3460 | 3480 | 3140 | 3273.73 | 3252.6244 | 2629 | 3179.8 | 3380 | 3430.9666 | 3422 | 3455.714 | 3192 | 2956 | 3123 | > 1000 | 3450 | 3285 | 1
| 1 | 2 | 2590 | 3116 | 3568.2 | 3270 | 3430 | 3620 | 3470 | 3420 | 3600 | 3120 | 3332.57 | 3325.6378 | 2555 | 3159.5 | 3340 | 3456.333 | 3465 | 3517.978 | 3157 | 2959 | 3077 | > 1000 | 3542 | 3244 | 17
| 1 | 3 | 2500 | 3094 | 3591.1 | 3220 | 3430 | 3650 | 3480 | 3440 | 3570 | 3100 | 3313.45 | 3311.2064 | 2533 | 3280.6 | 3340 | 3515.4741 | 3443 | 3541.838 | 3121 | 2927 | 3292 | > 1000 | 3420 | 3103 | 16
| 1 | 4 | 2410 | 3101 | 3663.3 | 3160 | 3380 | 3680 | 3460 | 3510 | 3530 | 3120 | 3436.69 | 3311.2064 | 2546 | 3331.8 | 3370 | 3501.0329 | 3513 | 3513.963 | 3142 | 2919 | 3210 | > 1000 | 3538 | 3162 | 3340.4787
| 1 | 5 | 2560 | 3092 | 3685.1 | 3090 | 3430 | 3730 | 3500 | 3470 | 3520 | 3160 | 3287.69 | 3240.2822 | 2533 | 3260.4 | 3330 | 3511.7399 | 3411 | 3585.567 | 3095 | 2885 | 3197 | > 1000 | 3422 | 3081 | 71
| 1 | 6 | 2360 | 3173 | 3652.1 | 3170 | 3470 | 3330 | 3500 | 3390 | 3460 | 3100 | 3186.21 | 3332.0499 | 2502 | 3240.4 | 3390 | 3403.5146 | 3457 | 3560.248 | 3126 | 2987 | 3212 | > 1000 | 3484 | 3160
| Mean |  | 2500 | 3092.3333 | 3630.9833 | 3171.6667 | 3421.6667 | 3608.3333 | 3485 | 3448.3333 | 3526.6667 | 3123.3333 | 3305.0567 | 3295.5012 | 2549.6667 | 3242.0833 | 3358.3333 | 3469.8435 | 3451.8333 | 3529.218 | 3138.8333 | 2938.8333 | 3185.1667 |  | 3476 | 3172.5
| Median |  | 2530 | 3097.5 | 3639.1 | 3165 | 3430 | 3645 | 3490 | 3450 | 3525 | 3120 | 3300.57 | 3311.2064 | 2539.5 | 3250.4 | 3355 | 3478.683 | 3450 | 3529.908 | 3134 | 2941.5 | 3203.5 |  | 3467 | 3161
| Std Dev. |  | 95.7079 | 63.5883 | 44.6658 | 65.549 | 32.5064 | 141.6216 | 17.6068 | 41.6733 | 52.7889 | 23.3809 | 81.9296 | 39.0509 | 42.8097 | 64.1407 | 24.8328 | 46.6984 | 36.268 | 44.8536 | 33.3791 | 35.9245 | 75.4385 |  | 54.7576 | 79.0411
| Rel.Std.Dev. |  | 3.828% | 2.056% | 1.23% | 2.067% | 0.95% | 3.925% | 0.5052% | 1.209% | 1.497% | 0.7486% | 2.479% | 1.185% | 1.679% | 1.978% | 0.7394% | 1.346% | 1.051% | 1.271% | 1.063% | 1.222% | 2.368% |  | 1.575% | 2.491%
| PDM3 |  | -25.16% | -7.428% | 8.696% | -5.054% | 2.43% | 8.018% | 4.326% | 3.229% | 5.574% | -6.5% | -1.06% | -1.346% | -23.67% | -2.946% | 0.5345% | 3.873% | 3.333% | 5.65% | -6.036% | -12.02% | -4.649% |  | 4.057% | -5.029%
| Z-Score (Absolute) |  | 3.27 | 0.86 | 1.32 | 0.54 | 0.47 | 1.23 | 0.73 | 0.58 | 0.9 | 0.74 | 0 | 0.04 | 3.06 | 0.26 | 0.22 | 0.67 | 0.6 | 0.91 | 0.67 | 1.49 | 0.49 | NA | 0.69 | 0.54
| Analytical results for B in OREAS 611 (Indicative Value < 10 ppm) | [25 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.18 | 2.20 | 2.24 | [8 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | [8 empty] | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | [8 empty] | 1
| 1 | 1 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 1
| 1 | 2 | < 10 | 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 8
| 1 | 3 | < 10 | 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 16
| 1 | 4 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | < 10
| 1 | 5 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15 | [8 empty] | 14
| 1 | 6 | < 10 | < 1 | < 1 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | < 10 | < 20 | < 20 | < 5 | < 10 | < 10 | 15
| Mean |  | 2.03: 1; 2.24: 15
| Median |  | 2.03: 1; 2.24: 15
| Std Dev. |  | 2.03: 0; 2.24: 0
Rel.Std.Dev. | bar: 2.03; 2.24
PDM3
Z-Score (Absolute) | bar: 2.02 to 2.24
| Analytical results for Ba in OREAS 611 (Indicative Value 58 ppm) | [25 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  | 0
| 1 | 1 | 253 | 21.9 | 11.4 | 30 | 30 | 20 | 20 | 30 | 224 | 12.55 | 71.6444 | 122 | 23.3 | 119.4984 | 272 | 101.629 | 133 | 96 | 138 | 58 | 20 | 2888 |  |  | 1
| 1 | 2 | 233 | 22.8 | 16.1 | 40 | 30 | 30 | 20 | 20 | 230 | 11.72 | 60.9639 | 106 | 23.3 | 111.3903 | 259 | 101.009 | 140 | 97 | 129 | 64 | 20 | 2960 |  |  | 5
| 1 | 3 | 272 | 22.7 | 13.9 | 30 | 20 | 20 | 20 | 30 | 227 | 11.42 | 65.1485 | 93 | 21.6 | 107.6805 | 255 | 106.669 | 130 | 127 | 140 | 65 | 20 | 3032 |  |  | 16
| 1 | 4 | 237 | 24.9 | 15.8 | 30 | 30 | 30 | 20 | 30 | 238 | 11.82 | 61.9362 | 89 | 22 | 117.1506 | 276 | 106.204 | 132 | 135 | 138 | 63 | 20 | 2987 |  |  | 57.7586
| 1 | 5 | 233 | 23.8 | 14 | 30 | 20 | 30 | 20 | 20 | 261 | 11.68 | 71.0005 | 100 | 22.1 | 118.1155 | 268 | 96.7 | 132 | 88 | 143 | 66 | 20 | 2945 |  |  | 15
| 1 | 6 | 156 | 23.1 | 15.7 | 20 | 30 | 20 | 20 | 30 | 260 | 12.43 | 64.575 | 80 | 23.1 | 118.2166 | 275 | 97.279 | 131 | 92 | 127 | 45 | 20 | 3021
| Mean |  | 230.6667 | 23.2 | 14.4833 | 30 | 26.6667 | 25 | 20 | 26.6667 | 240 | 11.9367 | 65.8781 | 98.3333 | 22.5667 | 115.342 | 267.5 | 101.5817 | 133 | 105.8333 | 135.8333 | 60.1667 | 20 | 2972.1667
| Median |  | 235 | 22.95 | 14.85 | 30 | 30 | 25 | 20 | 30 | 234 | 11.77 | 64.8618 | 96.5 | 22.6 | 117.6331 | 270 | 101.319 | 132 | 96.5 | 138 | 63.5 | 20 | 2973.5
| Std Dev. |  | 39.5811 | 1.0354 | 1.7837 | 6.3246 | 5.164 | 5.4772 | 0 | 5.164 | 16.5529 | 0.4501 | 4.5035 | 14.6515 | 0.75277 | 4.7078 | 8.6891 | 4.2408 | 3.5777 | 19.914 | 6.3692 | 7.9352 | 0 | 53.2068
| Rel.Std.Dev. |  | 17.16% | 4.463% | 12.32% | 21.08% | 19.36% | 21.91% | 0% | 19.36% | 6.897% | 3.771% | 6.836% | 14.9% | 3.336% | 4.082% | 3.248% | 4.175% | 2.69% | 18.82% | 4.689% | 13.19% | 0% | 1.79%
| PDM3 |  | 299.4% | -59.83% | -74.92% | -48.06% | -53.83% | -56.72% | -65.37% | -53.83% | 315.5% | -79.33% | 14.06% | 70.25% | -60.93% | 99.7% | 363.1% | 75.87% | 130.3% | 83.23% | 135.2% | 4.169% | -65.37% | 5046%
| Z-Score (Absolute) |  | 2.63 | 0.63 | 0.76 | 0.52 | 0.57 | 0.6 | 0.68 | 0.57 | 2.78 | 0.8 | 0.04 | 0.55 | 0.64 | 0.82 | 3.21 | 0.61 | 1.1 | 0.67 | 1.14 | 0.04 | 0.68 | 45.71
| Analytical results for Be in OREAS 611 (Certified Value 0.28 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 01g | [5 empty] | 3
| 1 | 1 | < 0.5 | 0.2 | 0.3 | 0.32 | 0.29 | 0.27 | 0.28 | 0.27 | < 0.5 | 0.29 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.3 | 0.2 | < 0.5 | < 1 | < 1 | [5 empty] | 1
| 1 | 2 | < 0.5 | 0.4 | 0.4 | 0.3 | 0.3 | 0.27 | 0.28 | 0.31 | < 0.5 | 0.28 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.4 | 0.2 | < 0.5 | < 1 | < 1 | [5 empty] | 32
| 1 | 3 | < 0.5 | 0.2 | 0.4 | 0.24 | 0.29 | 0.26 | 0.27 | 0.28 | < 0.5 | 0.28 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.3 | 0.3 | < 0.5 | < 1 | < 1 | [5 empty] | 16
| 1 | 4 | < 0.5 | 0.3 | 0.4 | 0.32 | 0.26 | 0.28 | 0.26 | 0.3 | < 0.5 | 0.28 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.4 | 0.2 | < 0.5 | < 1 | < 1 | [5 empty] | 0.28063
| 1 | 5 | < 0.5 | 0.4 | 0.3 | 0.3 | 0.3 | 0.31 | 0.27 | 0.28 | < 0.5 | 0.28 | < 0.5 | 0.29 | < 0.5 | < 0.5 | 0.4 | 0.2 | < 0.5 | < 1 | < 1 | [5 empty] | 72
| 1 | 6 | < 0.5 | 0.3 | 0.4 | 0.28 | 0.31 | 0.31 | 0.26 | 0.27 | < 0.5 | 0.28 | < 0.5 | 0.28 | < 0.5 | < 0.5 | 0.3 | 0.3 | < 0.5 | < 1 | < 1
| Mean |  | 2.03: 0.3; 2.04: 0.36667; 2.05: 0.29333; 2.06: 0.29167; 2.07: 0.28333; 2.08: 0.27; 2.09: 0.285; 2.11: 0.28167; 2.15: 0.29; 2.20: 0.35; 2.21: 0.23333
| Median |  | 2.03: 0.3; 2.04: 0.4; 2.05: 0.3; 2.06: 0.295; 2.07: 0.275; 2.08: 0.27; 2.09: 0.28; 2.11: 0.28; 2.15: 0.29; 2.20: 0.35; 2.21: 0.2
| Std Dev. |  | 2.03: 0.089443; 2.04: 0.05164; 2.05: 0.030111; 2.06: 0.017224; 2.07: 0.021602; 2.08: 0.0089443; 2.09: 0.016432; 2.11: 0.0040825; 2.15: 0.0063246; 2.20: 0.054772; 2.21: 0.05164
| Rel.Std.Dev. |  | 2.03: 29.81%; 2.04: 14.08%; 2.05: 10.27%; 2.06: 5.905%; 2.07: 7.624%; 2.08: 3.313%; 2.09: 5.766%; 2.11: 1.449%; 2.15: 2.181%; 2.20: 15.65%; 2.21: 22.13%
| PDM3 |  | 2.03: 6.904%; 2.04: 30.66%; 2.05: 4.529%; 2.06: 3.935%; 2.07: 0.9651%; 2.08: -3.786%; 2.09: 1.559%; 2.11: 0.3712%; 2.15: 3.341%; 2.20: 24.72%; 2.21: -16.85%
| Z-Score (Absolute) |  | 0.67 | 0.34 | 1.69 | 0.2 | 0.17 | 0 | 0.27 | 0.03 | 0.67 | 0.03 | 0.67 | 0.13 | 0.67 | 0.67 | 1.35 | 1.01 | 0.67 | 4.38 | 4.38
| Analytical results for Bi in OREAS 611 (Certified Value 256 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 231.1 | 279.69 | 222 | 260 | 277 | 264 | 260 | 267 | 259 | 267.76 | 238.5102 | 225 | 238.4 | 270 | 248.7752 | 247.53 | 266.059 | 257.23 | 251 | 247 | 260.2 | 375 | 164.8 |  | 1
| 1 | 2 | 240.42 | 277.44 | 230 | 261 | 275 | 261 | 260 | 270 | 264 | 260.87 | 239.1393 | 224 | 237.91 | 275 | 264.7828 | 248.45 | 270.921 | 257.21 | 246 | 244 | 261.3 | 379 | 147.5 |  | 33
| 1 | 3 | 240.35 | 277.78 | 233 | 260 | 279 | 264 | 259 | 270 | 260 | 260.44 | 243.4819 | 224 | 245.59 | 276 | 252.0614 | 238.83 | 273.554 | 263.22 | 250 | 263 | 266.5 | 385 | 150.2 |  | 16
| 1 | 4 | 241.13 | 291.77 | 223 | 260 | 280 | 266 | 266 | 266 | 264 | 263.49 | 243.1869 | 223 | 248.1 | 275 | 273.9673 | 246.6 | 271.202 | 248.71 | 251 | 260 | 268.7 | 384 | 152.1 |  | 256.488
| 1 | 5 | 237.09 | 288.34 | 225 | 260 | 284 | 269 | 265 | 272 | 256 | 259.11 | 241.789 | 220 | 242.54 | 277 | 273.6155 | 239.68 | 275.372 | 249.49 | 252 | 256 | 268 | 372 | 149.2 |  | 73
| 1 | 6 | 242.99 | 280.29 | 223 | 266 | 266 | 268 | 256 | 266 | 257 | 253.28 | 239.5713 | 216 | 243.72 | 274 | 251.0917 | 239.63 | 275.532 | 261 | 247 | 257 | 270.5 | 377 | 144.1
| Mean |  | 238.8467 | 282.5517 | 226 | 261.1667 | 276.8333 | 265.3333 | 261 | 268.5 | 260 | 260.825 | 240.9464 | 222 | 242.71 | 274.5 | 260.7157 | 243.4533 | 272.1067 | 256.1433 | 249.5 | 254.5 | 265.8667 | 378.6667 | 151.3167
| Median |  | 240.385 | 279.99 | 224 | 260 | 278 | 265 | 260 | 268.5 | 259.5 | 260.655 | 240.6801 | 223.5 | 243.13 | 275 | 258.4221 | 243.14 | 272.378 | 257.22 | 250.5 | 256.5 | 267.25 | 378 | 149.7
| Std Dev. |  | 4.2477 | 6.0114 | 4.4721 | 2.4014 | 6.1128 | 2.9439 | 3.7947 | 2.51 | 3.4059 | 4.8003 | 2.1564 | 3.4059 | 4.0003 | 2.429 | 11.5641 | 4.5104 | 3.5578 | 5.9246 | 2.429 | 7.4498 | 4.1812 | 5.0859 | 7.1385
| Rel.Std.Dev. |  | 1.778% | 2.128% | 1.979% | 0.9195% | 2.208% | 1.11% | 1.454% | 0.9348% | 1.31% | 1.84% | 0.895% | 1.534% | 1.648% | 0.8849% | 4.436% | 1.853% | 1.307% | 2.313% | 0.9735% | 2.927% | 1.573% | 1.343% | 4.718%
| PDM3 |  | -6.878% | 10.16% | -11.89% | 1.824% | 7.932% | 3.449% | 1.759% | 4.683% | 1.369% | 1.691% | -6.059% | -13.45% | -5.372% | 7.023% | 1.648% | -5.082% | 6.089% | -0.1344% | -2.724% | -0.7751% | 3.657% | 47.64% | -41%
| Z-Score (Absolute) |  | 1.29 | 1.29 | 2.06 | 0.03 | 0.95 | 0.27 | 0.02 | 0.46 | 0.04 | 0.01 | 1.17 | 2.29 | 1.07 | 0.82 | 0 | 1.02 | 0.67 | 0.27 | 0.66 | 0.37 | 0.3 | 6.98 | 6.48
| Analytical results for Ca in OREAS 611 (Certified Value 0.108 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.11 | 0.11 | 0.1 | 0.1 | 0.11 | 0.11 | 0.1 | 0.11 | 0.11 | 0.086 | 0.10465 | 0.09 | 0.1 | 0.11 | 0.10866 | 0.1 | 0.104 | 0.12 | 0.1 | 0.1 | 0.11 | 0.09 | 0.2852 |  | 1
| 1 | 2 | 0.11 | 0.1 | 0.11 | 0.1 | 0.11 | 0.11 | 0.1 | 0.11 | 0.1 | 0.097 | 0.10642 | 0.09 | 0.1 | 0.1 | 0.10836 | 0.1 | 0.104 | 0.11 | 0.1 | 0.1 | 0.11 | 0.09 | 0.3287 |  | #N/A
| 1 | 3 | 0.11 | 0.11 | 0.11 | 0.1 | 0.11 | 0.11 | 0.1 | 0.11 | 0.11 | 0.098 | 0.10479 | 0.09 | 0.11 | 0.11 | 0.10933 | 0.1 | 0.105 | 0.11 | 0.11 | 0.1 | 0.11 | 0.09 | 0.3122 |  | 16
| 1 | 4 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.1 | 0.10566 | 0.08 | 0.11 | 0.11 | 0.10943 | 0.1 | 0.104 | 0.12 | 0.1 | 0.1 | 0.11 | 0.09 | 0.3157 |  | 0.10798
| 1 | 5 | 0.12 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.11 | 0.095 | 0.10694 | 0.09 | 0.1 | 0.11 | 0.10921 | 0.1 | 0.104 | 0.1 | 0.11 | 0.1 | 0.11 | 0.09 | 0.3392 |  | 74
| 1 | 6 | 0.12 | 0.11 | 0.12 | 0.11 | 0.11 | 0.11 | 0.1 | 0.11 | 0.11 | 0.084 | 0.10636 | 0.08 | 0.1 | 0.11 | 0.10485 | 0.1 | 0.108 | 0.11 | 0.11 | 0.1 | 0.11 | 0.09 | 0.3127
| Mean |  | 0.11333 | 0.10833 | 0.11 | 0.105 | 0.11 | 0.11 | 0.10333 | 0.11 | 0.10833 | 0.093333 | 0.1058 | 0.086667 | 0.10333 | 0.10833 | 0.10831 | 0.1 | 0.10483 | 0.11167 | 0.105 | 0.1 | 0.11 | 0.09 | 0.31562
| Median |  | 0.11 | 0.11 | 0.11 | 0.105 | 0.11 | 0.11 | 0.1 | 0.11 | 0.11 | 0.096 | 0.10601 | 0.09 | 0.1 | 0.11 | 0.10894 | 0.1 | 0.104 | 0.11 | 0.105 | 0.1 | 0.11 | 0.09 | 0.3142
| Std Dev. |  | 0.005164 | 0.0040825 | 0.0063246 | 0.0054772 | 0 | 0 | 0.005164 | 0 | 0.0040825 | 0.0066833 | 0.00093649 | 0.005164 | 0.005164 | 0.0040825 | 0.0017434 | 1.5202e-17 | 0.0016021 | 0.0075277 | 0.0054772 | 1.5202e-17 | 0 | 1.5202e-17 | 0.01829
| Rel.Std.Dev. |  | 4.556% | 3.768% | 5.75% | 5.216% | 0% | 0% | 4.997% | 0% | 3.768% | 7.161% | 0.8851% | 5.958% | 4.997% | 3.768% | 1.61% | 1.52e-14% | 1.528% | 6.741% | 5.216% | 1.52e-14% | 0% | 1.689e-14% | 5.795%
| PDM3 |  | 4.957% | 0.3267% | 1.87% | -2.76% | 1.87% | 1.87% | -4.304% | 1.87% | 0.3267% | -13.56% | -2.016% | -19.74% | -4.304% | 0.3267% | 0.3036% | -7.391% | -2.915% | 3.414% | -2.76% | -7.391% | 1.87% | -16.65% | 192.3%
| Z-Score (Absolute) |  | 1.01 | 0 | 0.34 | 0.67 | 0.34 | 0.34 | 1.01 | 0.34 | 0 | 3.03 | 0.51 | 4.38 | 1.01 | 0 | 0.01 | NA | 0.71 | 0.67 | 0.67 | NA | 0.34 | 3.71 | 41.93
Results from laboratories 18 & 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cd in OREAS 611 (Certified Value 13.7 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 2
| 1 | 1 | 11.77 | 15.32 | 12.4 | 13.5 | 13.55 | 13.65 | 14.15 | 13.45 | 12.8 | 13.398 | 14.1354 | 10.4 | 13.3 | 14.8 | 13.8058 | 12.18 | 14.123 | 14.53 | 12.71 | 13 | 16.3 | 35 | 13.94 |  | 1
| 1 | 2 | 12.29 | 15.24 | 13 | 13.35 | 13.65 | 13.35 | 14.3 | 13.8 | 12.9 | 13.043 | 14.35 | 10.5 | 14.02 | 14 | 13.8578 | 11.92 | 14.062 | 14.42 | 13.22 | 13 | 16.1 | 35 | 14.67 |  | 34
| 1 | 3 | 11.92 | 15.5 | 13.3 | 13.4 | 13.65 | 13.55 | 14.75 | 13.4 | 12.9 | 13.046 | 14.7091 | 10.3 | 15.21 | 14.1 | 14.0337 | 12.03 | 14.274 | 14.28 | 13.02 | 14 | 16.3 | 35 | 15.4 |  | 16
| 1 | 4 | 12.3 | 16.22 | 13.2 | 13.5 | 13.75 | 13.5 | 13.8 | 13.4 | 12.6 | 13.543 | 13.802 | 10.5 | 15.11 | 14.6 | 13.9441 | 11.95 | 14.019 | 14.32 | 12.94 | 14 | 16.3 | 35 | 14.13 |  | 13.7293
| 1 | 5 | 12.06 | 16.25 | 12.9 | 13.7 | 14.05 | 13.9 | 14.4 | 13.15 | 12.7 | 13.143 | 14.1432 | 10.3 | 14.34 | 14.8 | 13.9501 | 11.77 | 14.333 | 13.98 | 13 | 13 | 16.6 | 35 | 15.86 |  | 75
| 1 | 6 | 12.42 | 15.54 | 13.2 | 13.95 | 13.05 | 13.75 | 13.25 | 13.4 | 12.9 | 13.219 | 14.4379 | 10 | 15.04 | 14.3 | 13.5404 | 12.41 | 15.612 | 14.26 | 13.24 | 13 | 14.7 | 35 | 16.59
| Mean |  | 12.1267 | 15.6783 | 13 | 13.5667 | 13.6167 | 13.6167 | 14.1083 | 13.4333 | 12.8 | 13.232 | 14.2629 | 10.3333 | 14.5033 | 14.4333 | 13.8553 | 12.0433 | 14.4038 | 14.2983 | 13.0217 | 13.3333 | 16.05 | 35 | 15.0983
| Median |  | 12.175 | 15.52 | 13.1 | 13.5 | 13.65 | 13.6 | 14.225 | 13.4 | 12.85 | 13.181 | 14.2466 | 10.35 | 14.69 | 14.45 | 13.9009 | 11.99 | 14.1985 | 14.3 | 13.01 | 13 | 16.3 | 35 | 15.035
| Std Dev. |  | 0.25185 | 0.44535 | 0.32863 | 0.22286 | 0.3266 | 0.19408 | 0.52289 | 0.20897 | 0.12649 | 0.20134 | 0.30995 | 0.18619 | 0.7568 | 0.35024 | 0.17339 | 0.22447 | 0.60421 | 0.18552 | 0.19559 | 0.5164 | 0.68044 | 0 | 1.0353
| Rel.Std.Dev. |  | 2.077% | 2.841% | 2.528% | 1.643% | 2.399% | 1.425% | 3.706% | 1.556% | 0.9882% | 1.522% | 2.173% | 1.802% | 5.218% | 2.427% | 1.251% | 1.864% | 4.195% | 1.297% | 1.502% | 3.873% | 4.24% | 0% | 6.857%
| PDM3 |  | -11.67% | 14.2% | -5.312% | -1.185% | -0.8207% | -0.8207% | 2.76% | -2.156% | -6.769% | -3.622% | 3.887% | -24.74% | 5.638% | 5.128% | 0.9176% | -12.28% | 4.913% | 4.144% | -5.154% | -2.884% | 16.9% | 154.9% | 9.971%
| Z-Score (Absolute) |  | 1.47 | 2.04 | 0.61 | 0.05 | 0 | 0 | 0.49 | 0.18 | 0.81 | 0.38 | 0.64 | 3.25 | 0.88 | 0.81 | 0.24 | 1.56 | 0.78 | 0.67 | 0.59 | NA | 2.41 | NA | 1.47
Results from laboratories 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ce in OREAS 611 (Certified Value 14.6 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 1
|  |  | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 2
| 1 | 1 | 14.4 | 17 | 13.4 | 13.45 | 14.45 | 14.95 | 14.95 | 24.5 | 15.641 | 21.4247 | 11.9 | 10.914 | 13.3 | 16.43 | 14.64 | 1 | [8 empty] | 1
| 1 | 2 | 13.8 | 18.7 | 13.9 | 14.05 | 14.6 | 15.15 | 15 | 24.1 | 15.001 | 20.8123 | 12.3 | 10.492 | 13.5 | 15.42 | 14.18 | 1 | [8 empty] | 35
| 1 | 3 | 14.4 | 18.5 | 13.65 | 13.75 | 14.7 | 15.35 | 14.4 | 24.5 | 14.839 | 20.8532 | 12.3 | 12.044 | 13 | 16.22 | 15.16 | 1 | [8 empty] | 16
| 1 | 4 | 14.6 | 18.2 | 14.1 | 14.45 | 14.4 | 14.8 | 14.3 | 24.5 | 15.138 | 20.9458 | 12.5 | 10.8561 | 13.7 | 15.59 | 14.95 | 1 | [8 empty] | 14.6378
| 1 | 5 | 15.5 | 17.8 | 14.65 | 13.95 | 15.6 | 15.3 | 13.95 | 24.3 | 15.014 | 21.6211 | 11.9 | 11.2906 | 12.9 | 15.44 | 15.13 | 1 | [8 empty] | 76
| 1 | 6 | 14.2 | 17.9 | 13.35 | 14.7 | 15.1 | 14.3 | 14.35 | 24.4 | 14.813 | 21.3281 | 11.8 | 10.2933 | 13.5 | 15.73 | 15.03 | < 1
| Mean |  | 14.4833 | 18.0167 | 13.8417 | 14.0583 | 14.8083 | 14.975 | 14.4917 | 24.3833 | 15.0743 | 21.1642 | 12.1167 | 10.9817 | 13.3167 | 15.805 | 14.8483 | 1
| Median |  | 14.4 | 18.05 | 13.775 | 14 | 14.65 | 15.05 | 14.375 | 24.45 | 15.0075 | 21.1369 | 12.1 | 10.8851 | 13.4 | 15.66 | 14.99 | 1
| Std Dev. |  | 0.56716 | 0.6047 | 0.4893 | 0.45653 | 0.46089 | 0.39083 | 0.40671 | 0.16021 | 0.30261 | 0.33816 | 0.28577 | 0.62578 | 0.31252 | 0.42336 | 0.3768 | 0
| Rel.Std.Dev. |  | 3.916% | 3.356% | 3.535% | 3.247% | 3.112% | 2.61% | 2.807% | 0.657% | 2.007% | 1.598% | 2.359% | 5.698% | 2.347% | 2.679% | 2.538% | 0%
| PDM3 |  | -1.055% | 23.08% | -5.439% | -3.958% | 1.165% | 2.304% | -0.998% | 66.58% | 2.983% | 44.59% | -17.22% | -24.98% | -9.025% | 7.974% | 1.439% | -93.17%
| Z-Score (Absolute) |  | 0.23 | 2.24 | 0.67 | 0.52 | 0 | 0.12 | 0.22 | 6.68 | 0.19 | 4.43 | 1.88 | 2.67 | 1.04 | 0.7 | 0.03 | NA
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 611 (Certified Value 8.64 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 2
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 3
| 1 | 1 | 6.58 | 8.6 | 9.9 | 8.8 | 8 | 8.6 | 9 | 8.6 | 7.9 | 9 | 8.5 | 7.7938 | 9 | 7.6 | 9.5 | 8.7561 | 8.36 | 8.4 | 8.7 | 8.6 | 8 | 9.5 | 8 | 4.602 | 1
| 1 | 2 | 6.2 | 8.9 | 9.8 | 8.9 | 8 | 8.7 | 8.8 | 8.9 | 8.3 | 8.9 | 8.63 | 8.9406 | 9 | 7.7 | 9.4 | 8.3999 | 8.47 | 8.618 | 8.7 | 8.3 | 8 | 9.5 | 8 | 4.487 | 21
| 1 | 3 | 6.23 | 8.8 | 10.1 | 9.1 | 8 | 8.9 | 8.8 | 8.8 | 8.1 | 9 | 8.59 | 8.9214 | 10 | 8.4 | 9.2 | 8.8795 | 8.11 | 8.649 | 8.6 | 8.5 | 8 | 9.5 | 8 | 4.372 | 16
| 1 | 4 | 6.28 | 9.1 | 10.3 | 8.5 | 8.6 | 8.9 | 8.8 | 8.7 | 7.8 | 9.1 | 9.09 | 8.4273 | 9 | 8.3 | 9.5 | 8.5322 | 8.35 | 8.673 | 8.4 | 8.2 | 8 | 9.5 | 8 | 4.257 | 8.6376
| 1 | 5 | 6.48 | 9 | 10.8 | 8.5 | 8.6 | 8.8 | 9 | 8.5 | 7.6 | 8.8 | 8.74 | 7.7663 | 10 | 8.3 | 9.5 | 8.494 | 8.15 | 8.515 | 8.2 | 8.1 | 8 | 9.8 | 8 | 4.142 | 77
| 1 | 6 | 6.42 | 9.1 | 9.8 | 8.5 | 8.2 | 8.5 | 8.9 | 8.4 | 8.2 | 9 | 8.28 | 8.7882 | 10 | 7.6 | 9.3 | 8.4058 | 8.3 | 8.732 | 8.5 | 8.4 | 8 | 9.8 | 8 | 4.027
| Mean |  | 6.365 | 8.9167 | 10.1167 | 8.7167 | 8.2333 | 8.7333 | 8.8833 | 8.65 | 7.9833 | 8.9667 | 8.6383 | 8.4396 | 9.5 | 7.9833 | 9.4 | 8.5779 | 8.29 | 8.5978 | 8.5167 | 8.35 | 8 | 9.6 | 8 | 4.3145
| Median |  | 6.35 | 8.95 | 10 | 8.65 | 8.1 | 8.75 | 8.85 | 8.65 | 8 | 9 | 8.61 | 8.6078 | 9.5 | 8 | 9.45 | 8.5131 | 8.325 | 8.6335 | 8.55 | 8.35 | 8 | 9.5 | 8 | 4.3145
| Std Dev. |  | 0.15176 | 0.19408 | 0.38687 | 0.25626 | 0.29439 | 0.1633 | 0.098319 | 0.18708 | 0.26394 | 0.10328 | 0.26992 | 0.54318 | 0.54772 | 0.38687 | 0.12649 | 0.19655 | 0.13638 | 0.12056 | 0.19408 | 0.18708 | 0 | 0.15492 | 0 | 0.21515
| Rel.Std.Dev. |  | 2.384% | 2.177% | 3.824% | 2.94% | 3.576% | 1.87% | 1.107% | 2.163% | 3.306% | 1.152% | 3.125% | 6.436% | 5.766% | 4.846% | 1.346% | 2.291% | 1.645% | 1.402% | 2.279% | 2.241% | 0% | 1.614% | 0% | 4.987%
| PDM3 |  | -26.31% | 3.231% | 17.12% | 0.9152% | -4.68% | 1.108% | 2.845% | 0.1434% | -7.575% | 3.81% | 0.008348% | -2.292% | 9.984% | -7.575% | 8.826% | -0.6911% | -4.024% | -0.4605% | -1.4% | -3.33% | -7.382% | 11.14% | -7.382% | -50.05%
| Z-Score (Absolute) |  | 4.89 | 0.7 | 3.33 | 0.26 | 0.8 | 0.3 | 0.63 | 0.11 | 1.35 | 0.81 | 0.09 | 0.35 | NA | 1.35 | 1.76 | 0.04 | 0.67 | 0 | 0.18 | 0.54 | NA | 2.2 | NA | 9.38
Results from laboratories 13, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 611 (Certified Value 41.7 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g |  | 2
| 1 | 1 | 13.4 | 46 | 45.3 | 43 | 40 | 43 | 40 | 41 | 41 | 45 | 37.9 | 44.4061 | 40 | 41.6 | 38 | 31.9752 | 41 | 40.048 | 44 | 41 | 38 | 42 | 36 |  | 1
| 1 | 2 | 12.4 | 49 | 46.2 | 45 | 40 | 44 | 41 | 41 | 42 | 45 | 39.2 | 41.5271 | 40 | 40.5 | 37.8 | 31.9294 | 41 | 40.585 | 43 | 40 | 37 | 42 | 36 |  | 37
| 1 | 3 | 12.5 | 49 | 47.1 | 45 | 40 | 44 | 41 | 40 | 42 | 44 | 39.3 | 43.4157 | 40 | 42.3 | 38.3 | 31.9036 | 40 | 40.48 | 42 | 40 | 40 | 43 | 35 |  | 16
| 1 | 4 | 14.2 | 49 | 47.5 | 43 | 40 | 45 | 41 | 42 | 41 | 44 | 41 | 41.2052 | 42 | 42.4 | 40.1 | 31.6069 | 42 | 40 | 43 | 39 | 39 | 43 | 36 |  | 41.6612
| 1 | 5 | 13.1 | 48 | 48.6 | 43 | 41 | 44 | 41 | 41 | 41 | 44 | 39.3 | 43.0172 | 41 | 41.4 | 38.2 | 32.2267 | 40 | 40.677 | 43 | 40 | 39 | 44 | 35 |  | 78
| 1 | 6 | 14.4 | 48 | 45.5 | 43 | 41 | 41 | 42 | 40 | 43 | 45 | 37.2 | 42.3873 | 41 | 41.7 | 38.7 | 30.8081 | 41 | 41.949 | 44 | 41 | 39 | 44 | 36
| Mean |  | 13.3333 | 48.1667 | 46.7 | 43.6667 | 40.3333 | 43.5 | 41 | 40.8333 | 41.6667 | 44.5 | 38.9833 | 42.6598 | 40.6667 | 41.65 | 38.5167 | 31.7416 | 40.8333 | 40.6232 | 43.1667 | 40.1667 | 38.6667 | 43 | 35.6667
| Median |  | 13.25 | 48.5 | 46.65 | 43 | 40 | 44 | 41 | 41 | 41.5 | 44.5 | 39.25 | 42.7022 | 40.5 | 41.65 | 38.25 | 31.9165 | 41 | 40.5325 | 43 | 40 | 39 | 43 | 36
| Std Dev. |  | 0.83825 | 1.169 | 1.2696 | 1.0328 | 0.5164 | 1.3784 | 0.63246 | 0.75277 | 0.8165 | 0.54772 | 1.3167 | 1.2018 | 0.8165 | 0.6892 | 0.83287 | 0.49821 | 0.75277 | 0.70737 | 0.75277 | 0.75277 | 1.0328 | 0.89443 | 0.5164
| Rel.Std.Dev. |  | 6.287% | 2.427% | 2.719% | 2.365% | 1.28% | 3.169% | 1.543% | 1.844% | 1.96% | 1.231% | 3.378% | 2.817% | 2.008% | 1.655% | 2.162% | 1.57% | 1.844% | 1.741% | 1.744% | 1.874% | 2.671% | 2.08% | 1.448%
| PDM3 |  | -68% | 15.62% | 12.09% | 4.814% | -3.187% | 4.414% | -1.587% | -1.987% | 0.01321% | 6.814% | -6.428% | 2.397% | -2.387% | -0.0268% | -7.548% | -23.81% | -1.987% | -2.492% | 3.614% | -3.587% | -7.188% | 3.214% | -14.39%
| Z-Score (Absolute) |  | 8.56 | 2.28 | 1.83 | 0.88 | 0.16 | 0.83 | 0.05 | 0 | 0.26 | 1.14 | 0.58 | 0.57 | 0.05 | 0.25 | 0.72 | 2.83 | 0 | 0.07 | 0.73 | 0.21 | 0.67 | 0.67 | 1.61
| Analytical results for Cs in OREAS 611 (Certified Value 0.73 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 0.84 | 0.75 | 0.78 | 0.66 | 0.67 | 0.72 | 0.71 | 0.71 | 0.89 | 0.73 | 0.91557 | 0.67 | 0.69 | 0.75 | 0.66 | 0.7 | [8 empty] | 1
| 1 | 2 | 0.87 | 0.78 | 0.78 | 0.7 | 0.68 | 0.71 | 0.71 | 0.71 | 0.89 | 0.71 | 0.93615 | 0.63 | 0.67 | 0.72 | 0.69 | 0.7 | [8 empty] | 38
| 1 | 3 | 0.88 | 0.76 | 0.79 | 0.68 | 0.7 | 0.72 | 0.74 | 0.69 | 0.87 | 0.73 | 0.93678 | 0.64 | 0.66 | 0.75 | 0.67 | 0.7 | [8 empty] | 16
| 1 | 4 | 0.88 | 0.8 | 0.77 | 0.67 | 0.72 | 0.7 | 0.72 | 0.69 | 0.87 | 0.74 | 0.88533 | 0.64 | 0.68 | 0.72 | 0.67 | 0.7 | [8 empty] | 0.73252
| 1 | 5 | 0.87 | 0.82 | 0.77 | 0.73 | 0.69 | 0.75 | 0.73 | 0.68 | 0.87 | 0.7 | 0.94218 | 0.64 | 0.66 | 0.71 | 0.67 | 0.7 | [8 empty] | 79
| 1 | 6 | 0.88 | 0.79 | 0.77 | 0.67 | 0.7 | 0.73 | 0.68 | 0.7 | 0.88 | 0.71 | 0.9012 | 0.63 | 0.69 | 0.71 | 0.67 | 0.7
| Mean |  | 0.87 | 0.78333 | 0.77667 | 0.685 | 0.69333 | 0.72167 | 0.715 | 0.69667 | 0.87833 | 0.72 | 0.91954 | 0.64167 | 0.675 | 0.72667 | 0.67167 | 0.7
| Median |  | 0.875 | 0.785 | 0.775 | 0.675 | 0.695 | 0.72 | 0.715 | 0.695 | 0.875 | 0.72 | 0.92586 | 0.64 | 0.675 | 0.72 | 0.67 | 0.7
| Std Dev. |  | 0.015492 | 0.02582 | 0.008165 | 0.025884 | 0.017512 | 0.017224 | 0.020736 | 0.012111 | 0.0098319 | 0.015492 | 0.022839 | 0.01472 | 0.013784 | 0.018619 | 0.0098319 | 1.2162e-16
| Rel.Std.Dev. |  | 1.781% | 3.296% | 1.051% | 3.779% | 2.526% | 2.387% | 2.9% | 1.738% | 1.119% | 2.152% | 2.484% | 2.294% | 2.042% | 2.562% | 1.464% | 1.737e-14%
| PDM3 |  | 18.77% | 6.936% | 6.026% | -6.488% | -5.35% | -1.482% | -2.392% | -4.895% | 19.91% | -1.71% | 25.53% | -12.4% | -7.853% | -0.7996% | -8.308% | -4.44%
| Z-Score (Absolute) |  | 2.25 | 0.95 | 0.85 | 0.52 | 0.4 | 0.02 | 0.07 | 0.35 | 2.37 | 0 | 2.99 | 1.17 | 0.67 | 0.1 | 0.72 | NA
Results from laboratory 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 611 (Certified Value 1.18 wt.%) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*AAS | AR*OES/MS | AR*OES/MS | AR*AAS | AR*OES | AR*AAS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*AAS | 2
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 0.5g | 01g | 0.5g | 01g | 01g | 0.5g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.2g | 01g | 3
| 1 | 1 | 1.18 | > 1 | > 1 | 1.16 | 1.115 | 1.15 | 1.18 | 1.155 | 1.165 | 1.18 | 1.1757 | 1.1703 | 1.209 | 1.15 | 1.17 | 1.1887 | 1.159 | 1.22 | 1.16 | 1.19 | > 1 | > 1 | 1.18 | 1.213 | 1
| 1 | 2 | 1.18 | > 1 | > 1 | 1.16 | 1.14 | 1.17 | 1.165 | 1.155 | 1.185 | 1.22 | 1.1769 | 1.1843 | 1.2 | 1.162 | 1.19 | 1.1896 | 1.156 | 1.16 | 1.17 | 1.19 | > 1 | > 1 | 1.19 | 1.19 | 23
| 1 | 3 | 1.21 | > 1 | > 1 | 1.17 | 1.14 | 1.155 | 1.18 | 1.16 | 1.16 | 1.2 | 1.1836 | 1.1796 | 1.202 | 1.155 | 1.18 | 1.2064 | 1.162 | 1.19 | 1.17 | 1.2 | > 1 | > 1 | 1.2 | 1.167 | 16
| 1 | 4 | 1.17 | > 1 | > 1 | 1.17 | 1.15 | 1.16 | 1.17 | 1.16 | 1.24 | 1.19 | 1.2037 | 1.1748 | 1.186 | 1.136 | 1.2 | 1.2032 | 1.16 | 1.19 | 1.16 | 1.22 | > 1 | > 1 | 1.21 | 1.144 | 1.1791
| 1 | 5 | 1.22 | > 1 | > 1 | 1.2 | 1.16 | 1.15 | 1.2 | 1.16 | 1.23 | 1.2 | 1.1685 | 1.1879 | 1.213 | 1.171 | 1.19 | 1.2067 | 1.152 | 1.19 | 1.16 | 1.2 | > 1 | > 1 | 1.19 | 1.201 | 80
| 1 | 6 | 1.23 | > 1 | > 1 | 1.19 | 1.13 | 1.145 | 1.195 | 1.16 | 1.185 | 1.22 | 1.1357 | 1.1792 | 1.197 | 1.154 | 1.17 | 1.173 | 1.159 | 1.23 | 1.18 | 1.18 | > 1 | > 1 | 1.17 | 1.298
| Mean |  | 1.1983 |  |  | 1.175 | 1.1392 | 1.155 | 1.1817 | 1.1583 | 1.1942 | 1.2017 | 1.174 | 1.1793 | 1.2012 | 1.1547 | 1.1833 | 1.1946 | 1.158 | 1.1967 | 1.1667 | 1.1967 |  |  | 1.19 | 1.2022
| Median |  | 1.195 |  |  | 1.17 | 1.14 | 1.1525 | 1.18 | 1.16 | 1.185 | 1.2 | 1.1763 | 1.1794 | 1.201 | 1.1545 | 1.185 | 1.1964 | 1.159 | 1.19 | 1.165 | 1.195 |  |  | 1.19 | 1.1955
| Std Dev. |  | 0.024833 |  |  | 0.016432 | 0.015626 | 0.0089443 | 0.013663 | 0.002582 | 0.033379 | 0.016021 | 0.022285 | 0.0063344 | 0.0094956 | 0.011759 | 0.012111 | 0.013324 | 0.0035214 | 0.025033 | 0.008165 | 0.013663 |  |  | 0.014142 | 0.053049
| Rel.Std.Dev. |  | 2.072% |  |  | 1.398% | 1.372% | 0.7744% | 1.156% | 0.2229% | 2.795% | 1.333% | 1.898% | 0.5371% | 0.7905% | 1.018% | 1.023% | 1.115% | 0.3041% | 2.092% | 0.6999% | 1.142% |  |  | 1.188% | 4.413%
| PDM3 |  | 1.634% |  |  | -0.3454% | -3.385% | -2.042% | 0.22% | -1.759% | 1.28% | 1.916% | -0.43% | 0.02296% | 1.874% | -2.07% | 0.3613% | 1.318% | -1.787% | 1.492% | -1.052% | 1.492% |  |  | 0.9268% | 1.959%
| Z-Score (Absolute) |  | 0.71 | NA | NA | 0.34 | 1.95 | 1.24 | 0.04 | 1.09 | 0.52 | 0.86 | 0.38 | 0.14 | 0.84 | 1.25 | 0.04 | 0.54 | 1.1 | 0.64 | 0.71 | 0.64 | NA | NA | 0.34 | 0.88
| Analytical results for Dy in OREAS 611 (Indicative Value 0.83 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 0.8 | 1
| 1 | 2 | 2.01: 0.9 | 10
| 1 | 3 | 2.01: 0.8 | 16
| 1 | 4 | 2.01: 0.8 | 0.83333
| 1 | 5 | 2.01: 0.8 | 16
| 1 | 6 | 2.01: 0.9
| Mean |  | 2.01: 0.83333
| Median |  | 2.01: 0.8
| Std Dev. |  | 2.01: 0.05164
| Rel.Std.Dev. |  | 2.01: 6.197%
| PDM3 |  | 2.01: 4.441e-14%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Er in OREAS 611 (Indicative Value 0.28 ppm) | [25 empty] | Indicative
Er | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 0.2 | 1
| 1 | 2 | 2.01: 0.3 | 11
| 1 | 3 | 2.01: 0.3 | 16
| 1 | 4 | 2.01: 0.3 | 0.28333
| 1 | 5 | 2.01: 0.3 | 17
| 1 | 6 | 2.01: 0.3
| Mean |  | 2.01: 0.28333
| Median |  | 2.01: 0.3
| Std Dev. |  | 2.01: 0.040825
| Rel.Std.Dev. |  | 2.01: 14.41%
| PDM3 |  | 2.01: 1.332e-13%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Eu in OREAS 611 (Indicative Value 0.22 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 0.2 | 1
| 1 | 2 | 2.01: 0.2 | 12
| 1 | 3 | 2.01: 0.3 | 16
| 1 | 4 | 2.01: 0.2 | 0.21667
| 1 | 5 | 2.01: 0.2 | 18
| 1 | 6 | 2.01: 0.2
| Mean |  | 2.01: 0.21667
| Median |  | 2.01: 0.2
| Std Dev. |  | 2.01: 0.040825
| Rel.Std.Dev. |  | 2.01: 18.84%
| PDM3 |  | 2.01: -1.665e-13%
Z-Score (Absolute) | bar: 2.01
| Analytical results for Fe in OREAS 611 (Certified Value 2.4 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | wt.%
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 2
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 3
| 1 | 1 | 2.32 | 2.57 | 2.43 | 2.41 | 2.28 | 2.43 | 2.33 | 2.3 | 2.31 | 2.76 | 2.255 | 2.5936 | 2.05 | 2.29 | 2.41 | 2.3475 | 2.47 | 2.494 | 2.43 | 2.52 | 2.38 | 2.15 | 2.32 | 2.697 | 1
| 1 | 2 | 2.31 | 2.66 | 2.39 | 2.48 | 2.29 | 2.43 | 2.34 | 2.29 | 2.35 | 2.73 | 2.363 | 2.5149 | 2.04 | 2.29 | 2.31 | 2.3404 | 2.46 | 2.388 | 2.43 | 2.51 | 2.51 | 2.14 | 2.37 | 2.634 | #N/A
| 1 | 3 | 2.29 | 2.64 | 2.42 | 2.4 | 2.28 | 2.45 | 2.34 | 2.28 | 2.34 | 2.74 | 2.364 | 2.5756 | 2.03 | 2.36 | 2.46 | 2.3414 | 2.41 | 2.463 | 2.42 | 2.56 | 2.52 | 2.13 | 2.34 | 2.671 | 16
| 1 | 4 | 2.27 | 2.67 | 2.43 | 2.33 | 2.29 | 2.47 | 2.33 | 2.33 | 2.32 | 2.77 | 2.421 | 2.5133 | 2.07 | 2.38 | 2.5 | 2.3525 | 2.48 | 2.41 | 2.42 | 2.54 | 2.48 | 2.13 | 2.35 | 2.6508 | 2.4047
| 1 | 5 | 2.31 | 2.63 | 2.45 | 2.42 | 2.3 | 2.51 | 2.38 | 2.31 | 2.34 | 2.78 | 2.344 | 2.6111 | 2.05 | 2.34 | 2.42 | 2.3549 | 2.43 | 2.471 | 2.42 | 2.56 | 2.46 | 2.14 | 2.31 | 2.445 | 81
| 1 | 6 | 2.3 | 2.67 | 2.43 | 2.33 | 2.31 | 2.31 | 2.37 | 2.23 | 2.29 | 2.74 | 2.193 | 2.5903 | 2.01 | 2.34 | 2.41 | 2.2793 | 2.5 | 2.564 | 2.45 | 2.57 | 2.43 | 2.14 | 2.31 | 2.582
| Mean |  | 2.3 | 2.64 | 2.425 | 2.395 | 2.2917 | 2.4333 | 2.3483 | 2.29 | 2.325 | 2.7533 | 2.3233 | 2.5665 | 2.0417 | 2.3333 | 2.4183 | 2.336 | 2.4583 | 2.465 | 2.4283 | 2.5433 | 2.4633 | 2.1383 | 2.3333 | 2.6133
| Median |  | 2.305 | 2.65 | 2.43 | 2.405 | 2.29 | 2.44 | 2.34 | 2.295 | 2.33 | 2.75 | 2.3535 | 2.583 | 2.045 | 2.34 | 2.415 | 2.3445 | 2.465 | 2.467 | 2.425 | 2.55 | 2.47 | 2.14 | 2.33 | 2.6424
| Std Dev. |  | 0.017889 | 0.037947 | 0.019748 | 0.057533 | 0.01169 | 0.067429 | 0.02137 | 0.034059 | 0.022583 | 0.019664 | 0.083486 | 0.042119 | 0.020412 | 0.036697 | 0.063692 | 0.02837 | 0.033116 | 0.062667 | 0.01169 | 0.024221 | 0.052409 | 0.0075277 | 0.024221 | 0.091076
| Rel.Std.Dev. |  | 0.7778% | 1.437% | 0.8144% | 2.402% | 0.5101% | 2.771% | 0.91% | 1.487% | 0.9713% | 0.7142% | 3.593% | 1.641% | 0.9998% | 1.573% | 2.634% | 1.214% | 1.347% | 2.542% | 0.4814% | 0.9523% | 2.128% | 0.352% | 1.038% | 3.485%
| PDM3 |  | -4.353% | 9.786% | 0.8454% | -0.4022% | -4.699% | 1.192% | -2.343% | -4.769% | -3.313% | 14.5% | -3.382% | 6.728% | -15.1% | -2.967% | 0.5682% | -2.856% | 2.232% | 2.509% | 0.984% | 5.766% | 2.44% | -11.08% | -2.967% | 8.676%
| Z-Score (Absolute) |  | 0.93 | 2.03 | 0.16 | 0.1 | 1 | 0.23 | 0.51 | 1.02 | 0.71 | 3.02 | 0.73 | 1.39 | 3.18 | 0.64 | 0.1 | 0.61 | 0.45 | 0.51 | 0.19 | 1.19 | 0.49 | 2.33 | 0.64 | 1.8
| Analytical results for Ga in OREAS 611 (Certified Value 6.83 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 01g |  |  |  | 3
| 1 | 1 | 7 | 7 | 8.4 | 6.17 | 6.61 | 6.85 | 6.44 | 6.66 | 6.6 | 6.6 | 7.7034 | < 5 | 7.5 | 6.65 | 6.1802 | 5.37 | 6.4 | 6.8 | 6 | 7 | 10.82 |  |  |  | 1
| 1 | 2 | 7.5 | 7 | 8.72 | 6.08 | 6.73 | 6.83 | 6.55 | 6.91 | 6.8 | 6.8 | 7.8567 | < 5 | 7.6 | 6.68 | 5.6508 | 5.37 | 6.4 | 6.8 | 6 | 7 | 10.38 |  |  |  | 43
| 1 | 3 | 7.8 | 7.3 | 8.68 | 6.3 | 6.68 | 6.86 | 6.69 | 6.78 | 6.6 | 6.7 | 7.9213 | < 5 | 7.6 | 6.53 | 6.2182 | 5.26 | 6.6 | 7.2 | 6 | 7 | 9.94 |  |  |  | 16
| 1 | 4 | 7.5 | 7.1 | 8.36 | 6.39 | 6.92 | 6.75 | 6.26 | 6.66 | 6.7 | 7 | 7.9091 | < 5 | 7.6 | 6.66 | 5.4877 | 5.47 | 6.2 | 6.7 | 6 | 7 | 10.5 |  |  |  | 6.83
| 1 | 5 | 7.4 | 7.4 | 8.18 | 6.47 | 6.71 | 6.78 | 6.43 | 6.66 | 6.8 | 6.7 | 8.0476 | < 5 | 7.5 | 6.83 | 6.4248 | 5.29 | 6.2 | 7 | 7 | 7 | 10.06 |  |  |  | 82
| 1 | 6 | 7.6 | 6.8 | 8.3 | 6.46 | 6.86 | 6.74 | 6.63 | 6.72 | 6.6 | 6.4 | 7.7304 | < 5 | 7.9 | 6.74 | 6.1572 | 5.43 | 6.6 | 7 | 6 | 7 | 9.62
| Mean |  | 7.4667 | 7.1 | 8.44 | 6.3117 | 6.7517 | 6.8017 | 6.5 | 6.7317 | 6.6833 | 6.7 | 7.8614 |  | 7.6167 | 6.6817 | 6.0198 | 5.365 | 6.4 | 6.9167 | 6.1667 | 7 | 10.22
| Median |  | 7.5 | 7.05 | 8.38 | 6.345 | 6.72 | 6.805 | 6.495 | 6.69 | 6.65 | 6.7 | 7.8829 |  | 7.6 | 6.67 | 6.1687 | 5.37 | 6.4 | 6.9 | 6 | 7 | 10.22
| Std Dev. |  | 0.26583 | 0.21909 | 0.21503 | 0.15943 | 0.11618 | 0.051929 | 0.15595 | 0.099683 | 0.098319 | 0.2 | 0.12861 |  | 0.1472 | 0.099883 | 0.36532 | 0.079937 | 0.17889 | 0.18348 | 0.40825 | 0 | 0.43081
| Rel.Std.Dev. |  | 3.56% | 3.086% | 2.548% | 2.526% | 1.721% | 0.7635% | 2.399% | 1.481% | 1.471% | 2.985% | 1.636% |  | 1.933% | 1.495% | 6.069% | 1.49% | 2.795% | 2.653% | 6.62% | 0% | 4.215%
| PDM3 |  | 9.321% | 3.953% | 23.57% | -7.59% | -1.147% | -0.4154% | -4.832% | -1.44% | -2.148% | -1.904% | 15.1% |  | 11.52% | -2.172% | -11.86% | -21.45% | -6.296% | 1.268% | -9.713% | 2.488% | 49.63%
| Z-Score (Absolute) |  | 1.35 | 0.67 | 3.13 | 0.77 | 0.04 | 0.13 | 0.42 | 0 | 0.09 | 0.06 | 2.07 | 7.75 | 1.62 | 0.09 | 1.3 | 2.5 | 0.61 | 0.34 | NA | NA | 6.39
Results from laboratories 23 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 611 (Indicative Value 1.47 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | 2.02: Lab |  |  |  | 1
| No | No | 2.02: 2.04 |  |  |  | ppm
|  |  | 2.02: AR*MS |  |  |  | 2
|  |  | 2.02: 0.5g |  |  |  | 2
| 1 | 1 | 2.02: 1.4 |  |  |  | 1
| 1 | 2 | 2.02: 1.4 |  |  |  | 13
| 1 | 3 | 2.02: 1.6 |  |  |  | 16
| 1 | 4 | 2.02: 1.5 |  |  |  | 1.4667
| 1 | 5 | 2.02: 1.4 |  |  |  | 19
| 1 | 6 | 2.02: 1.5
| Mean |  | 2.02: 1.4667
| Median |  | 2.02: 1.45
| Std Dev. |  | 2.02: 0.08165
| Rel.Std.Dev. |  | 2.02: 5.567%
| PDM3 |  | 2.02: -2.109e-13%
Z-Score (Absolute) | bar: 2.02
| Analytical results for Ge in OREAS 611 (Indicative Value 0.32 ppm) | [25 empty] | Indicative
Ge | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab; 2.06: Lab; 2.07: Lab; 2.08: Lab; 2.09: Lab; 2.10: Lab; 2.11: Lab; 2.12: Lab; 2.13: Lab |  |  |  | 1
| No | No | 2.02: 2.03; 2.03: 2.04; 2.04: 2.05; 2.05: 2.06; 2.06: 2.07; 2.07: 2.08; 2.08: 2.09; 2.09: 2.11; 2.10: 2.18; 2.11: 2.20; 2.12: 2.21; 2.13: 2.23 |  |  |  | ppm
|  |  | 2.02: AR*MS; 2.03: AR*MS; 2.04: AR*MS; 2.05: AR*MS; 2.06: AR*MS; 2.07: AR*MS; 2.08: AR*MS; 2.09: AR*MS; 2.10: AR*OES/MS; 2.11: AR*OES/MS; 2.12: AR*OES/MS; 2.13: AR*OES/MS |  |  |  | 2
|  |  | 2.02: 15g; 2.03: 0.5g; 2.04: 0.5g; 2.05: 0.5g; 2.06: 0.5g; 2.07: 0.5g; 2.08: 0.5g; 2.09: 0.5g; 2.10: 0.2g; 2.11: 0.25g; 2.12: 0.25g; 2.13: 0.25g |  |  |  | 2
| 1 | 1 | 2.02: 0.8; 2.03: 0.2; 2.04: 0.5; 2.05: 0.26; 2.06: 0.45; 2.07: 0.25; 2.08: 0.17; 2.09: 0.48; 2.10: 0.11; 2.11: 0.9; 2.12: 0.7; 2.13: < 0.1 |  |  |  | 1
| 1 | 2 | 2.02: 1; 2.03: 0.7; 2.04: 0.5; 2.05: 0.29; 2.06: 0.43; 2.07: 0.25; 2.08: 0.18; 2.09: 0.5; 2.10: 0.11; 2.11: 0.8; 2.12: 0.6; 2.13: < 0.1 |  |  |  | 6
| 1 | 3 | 2.02: 0.9; 2.03: 0.4; 2.04: 0.52; 2.05: 0.29; 2.06: 0.45; 2.07: 0.27; 2.08: 0.15; 2.09: 0.5; 2.10: 0.11; 2.11: 0.7; 2.12: 0.7; 2.13: < 0.1 |  |  |  | 16
| 1 | 4 | 2.02: 0.9; 2.03: 0.5; 2.04: 0.5; 2.05: 0.29; 2.06: 0.39; 2.07: 0.23; 2.08: 0.18; 2.09: 0.53; 2.10: 0.12; 2.11: 0.7; 2.12: 0.7; 2.13: < 0.1 |  |  |  | 0.32071
| 1 | 5 | 2.02: 1; 2.03: 0.4; 2.04: 0.54; 2.05: 0.32; 2.06: 0.4; 2.07: 0.25; 2.08: 0.17; 2.09: 0.45; 2.10: 0.11; 2.11: 0.8; 2.12: 0.8; 2.13: < 0.1 |  |  |  | 20
| 1 | 6 | 2.02: 0.9; 2.03: 0.5; 2.04: 0.55; 2.05: 0.24; 2.06: 0.39; 2.07: 0.34; 2.08: 0.18; 2.09: 0.5; 2.10: 0.11; 2.11: 0.9; 2.12: 0.8; 2.13: < 0.1
| Mean |  | 2.02: 0.91667; 2.03: 0.45; 2.04: 0.51833; 2.05: 0.28167; 2.06: 0.41833; 2.07: 0.265; 2.08: 0.17167; 2.09: 0.49333; 2.10: 0.11167; 2.11: 0.8; 2.12: 0.71667
| Median |  | 2.02: 0.9; 2.03: 0.45; 2.04: 0.51; 2.05: 0.29; 2.06: 0.415; 2.07: 0.25; 2.08: 0.175; 2.09: 0.5; 2.10: 0.11; 2.11: 0.8; 2.12: 0.7
| Std Dev. |  | 2.02: 0.075277; 2.03: 0.16432; 2.04: 0.022286; 2.05: 0.027869; 2.06: 0.028577; 2.07: 0.038859; 2.08: 0.01169; 2.09: 0.026583; 2.10: 0.0040825; 2.11: 0.089443; 2.12: 0.075277
| Rel.Std.Dev. |  | 2.02: 8.212%; 2.03: 36.51%; 2.04: 4.3%; 2.05: 9.894%; 2.06: 6.831%; 2.07: 14.66%; 2.08: 6.81%; 2.09: 5.388%; 2.10: 3.656%; 2.11: 11.18%; 2.12: 10.5%
| PDM3 |  | 2.02: 185.8%; 2.03: 40.31%; 2.04: 61.62%; 2.05: -12.18%; 2.06: 30.44%; 2.07: -17.37%; 2.08: -46.47%; 2.09: 53.82%; 2.10: -65.18%; 2.11: 149.4%; 2.12: 123.5%
| Z-Score (Absolute) |  | 2.02: NA; 2.03: NA; 2.04: 1.08; 2.05: 0.04; 2.06: 0.64; 2.07: 0.04; 2.08: 0.45; 2.09: 0.97; 2.10: 0.71; 2.11: NA; 2.12: NA; 2.13: 0.98
Results from laboratories 3, 4, 20 & 21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 611 (Certified Value 0.42 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 0.4 | 0.49 | 0.5 | 0.43 | 0.45 | 0.44 | 0.46 | 0.42 | 0.36 | 0.45 | 0.3857 | 0.4 | 0.39 | 0.43 | 0.39 | 0.5 | [8 empty] | 1
| 1 | 2 | 0.4 | 0.45 | 0.5 | 0.44 | 0.42 | 0.45 | 0.46 | 0.44 | 0.38 | 0.43 | 0.38417 | 0.41 | 0.39 | 0.4 | 0.39 | 0.5 | [8 empty] | 29
| 1 | 3 | 0.42 | 0.46 | 0.5 | 0.43 | 0.45 | 0.45 | 0.48 | 0.44 | 0.39 | 0.42 | 0.39543 | 0.4 | 0.38 | 0.44 | 0.38 | 0.5 | [8 empty] | 16
| 1 | 4 | 0.42 | 0.44 | 0.5 | 0.43 | 0.48 | 0.45 | 0.47 | 0.41 | 0.38 | 0.43 | 0.3984 | 0.4 | 0.39 | 0.4 | 0.38 | 0.6 | [8 empty] | 0.41875
| 1 | 5 | 0.4 | 0.46 | 0.5 | 0.42 | 0.47 | 0.43 | 0.44 | 0.42 | 0.39 | 0.43 | 0.3784 | 0.38 | 0.38 | 0.43 | 0.4 | 0.6 | [8 empty] | 83
| 1 | 6 | 0.41 | 0.45 | 0.5 | 0.45 | 0.43 | 0.44 | 0.47 | 0.44 | 0.38 | 0.41 | 0.38296 | 0.38 | 0.39 | 0.4 | 0.36 | 0.5
| Mean |  | 0.40833 | 0.45833 | 0.5 | 0.43333 | 0.45 | 0.44333 | 0.46333 | 0.42833 | 0.38 | 0.42833 | 0.38751 | 0.395 | 0.38667 | 0.41667 | 0.38333 | 0.53333
| Median |  | 0.405 | 0.455 | 0.5 | 0.43 | 0.45 | 0.445 | 0.465 | 0.43 | 0.38 | 0.43 | 0.38494 | 0.4 | 0.39 | 0.415 | 0.385 | 0.5
| Std Dev. |  | 0.0098319 | 0.017224 | 0 | 0.010328 | 0.022804 | 0.008165 | 0.013663 | 0.013292 | 0.010954 | 0.013292 | 0.0077359 | 0.012247 | 0.005164 | 0.018619 | 0.013663 | 0.05164
| Rel.Std.Dev. |  | 2.408% | 3.758% | 0% | 2.383% | 5.067% | 1.842% | 2.949% | 3.103% | 2.883% | 3.103% | 1.996% | 3.101% | 1.336% | 4.469% | 3.564% | 9.682%
| PDM3 |  | -2.488% | 9.453% | 19.4% | 3.482% | 7.463% | 5.87% | 10.65% | 2.288% | -9.254% | 2.288% | -7.46% | -5.672% | -7.662% | -0.4977% | -8.458% | 27.36%
| Z-Score (Absolute) |  | 0.35 | 0.88 | NA | 0.27 | 0.67 | 0.51 | 1 | 0.14 | 1.04 | 0.14 | 0.86 | 0.67 | 0.88 | 0.14 | 0.96 | NA
Results from laboratories 4 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 611 (Certified Value 0.95 ppm) | [25 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.11 | 2.12 | 2.13 | 2.14 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 10g | 0.25g | 01g | [6 empty] | 3
| 1 | 1 | 0.943 | 0.71 | 0.91 | 0.97 | 0.98 | 0.89 | 0.88 | 0.98 | 0.77452 | 1 | 0.957 | < 2 | < 1 | 0.97 | 0.9 | < 5 | 1.001 | < 5 | [6 empty] | 1
| 1 | 2 | 0.922 | 0.75 | 0.89 | 0.9 | 0.99 | 0.89 | 0.89 | 0.99 | 0.82133 | 1 | 0.957 | < 2 | < 1 | 0.96 | 0.92 | < 5 | 0.979 | < 5 | [6 empty] | 5
| 1 | 3 | 0.93 | 0.76 | 0.91 | 0.92 | 0.97 | 0.9 | 0.9 | 0.96 | 0.82004 | 1 | 1.116 | < 2 | < 1 | 0.97 | 0.94 | < 5 | 0.98 | < 5 | [6 empty] | 16
| 1 | 4 | 0.921 | 0.75 | 0.95 | 1.01 | 0.99 | 0.88 | 0.86 | 0.98 | 0.79339 | 2 | 1.086 | < 2 | < 1 | 0.92 | 0.86 | < 5 | 1.002 | < 5 | [6 empty] | 0.94658
| 1 | 5 | 1.006 | 0.73 | 0.92 | 0.99 | 0.93 | 0.89 | 0.9 | 0.96 | 0.85898 | < 1 | 0.992 | < 2 | < 1 | 0.95 | 0.94 | < 5 | 0.984 | < 5 | [6 empty] | 84
| 1 | 6 | 0.971 | 0.69 | 0.98 | 0.97 | 0.95 | 0.89 | 0.85 | 0.93 | 0.85499 | < 1 | 1.01 | < 2 | < 1 | 0.97 | 0.95 | < 5 | 0.857 | < 5
| Mean |  | 0.94883 | 0.73167 | 0.92667 | 0.96 | 0.96833 | 0.89 | 0.88 | 0.96667 | 0.82054 | 1.25 | 1.0197 |  |  | 0.95667 | 0.91833 |  | 0.96717
| Median |  | 0.9365 | 0.74 | 0.915 | 0.97 | 0.975 | 0.89 | 0.885 | 0.97 | 0.82069 | 1 | 1.001 |  |  | 0.965 | 0.93 |  | 0.982
| Std Dev. |  | 0.033594 | 0.027142 | 0.03266 | 0.041952 | 0.024014 | 0.0063246 | 0.020976 | 0.021602 | 0.033228 | 0.5 | 0.066923 |  |  | 0.019664 | 0.033714 |  | 0.054923
| Rel.Std.Dev. |  | 3.541% | 3.71% | 3.524% | 4.37% | 2.48% | 0.7106% | 2.384% | 2.235% | 4.05% | 40% | 6.563% |  |  | 2.055% | 3.671% |  | 5.679%
| PDM3 |  | 0.2375% | -22.7% | -2.104% | 1.417% | 2.298% | -5.978% | -7.034% | 2.122% | -13.32% | 32.05% | 7.721% |  |  | 1.065% | -2.985% |  | 2.174%
| Z-Score (Absolute) |  | 0.12 | 3.5 | 0.47 | 0.05 | 0.18 | 1.04 | 1.19 | 0.16 | 2.12 | NA | 0.98 | 0.67 | 7.11 | 0 | 0.6 | 24.02 | 0.16 | 24.02
| Analytical results for Ho in OREAS 611 (Indicative Value 0.1 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | 2.03: Lab | [6 empty] | 1
| No | No | 2.03: 2.04 | [6 empty] | ppm
|  |  | 2.03: AR*MS | [6 empty] | 2
|  |  | 2.03: 0.5g | [6 empty] | 2
| 1 | 1 | 2.03: 0.1 | [6 empty] | 1
| 1 | 2 | 2.03: 0.1 | [6 empty] | 8
| 1 | 3 | 2.03: 0.1 | [6 empty] | 16
| 1 | 4 | 2.03: 0.1 | [6 empty] | 0.1
| 1 | 5 | 2.03: 0.1 | [6 empty] | 14
| 1 | 6 | 2.03: 0.1
| Mean |  | 2.03: 0.1
| Median |  | 2.03: 0.1
| Std Dev. |  | 2.03: 1.5202e-17
| Rel.Std.Dev. |  | 2.03: 1.52e-14%
| PDM3 |  | 2.03: -1.11e-14%
Z-Score (Absolute) | bar: 2.03
| Analytical results for In in OREAS 611 (Certified Value 4.53 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 3.89 | 4.81 | 4.5 | 4.27 | 4.63 | 4.5 | 4.97 | 4.52 | 4.52 | 4.49 | 4.6097 | 4.63 | 6.3 | 4.66 | 4.37 | 4.8 | [8 empty] | 1
| 1 | 2 | 4.12 | 5.06 | 4.65 | 4.21 | 4.76 | 4.39 | 5.04 | 4.61 | 4.63 | 4.44 | 4.5759 | 4.54 | 6.2 | 4.58 | 4.07 | 4.8 | [8 empty] | 31
| 1 | 3 | 4.03 | 4.92 | 4.5 | 4.42 | 4.71 | 4.54 | 5.25 | 4.47 | 4.42 | 4.43 | 4.6169 | 4.4 | 6.2 | 4.61 | 4.31 | 4.8 | [8 empty] | 16
| 1 | 4 | 4.09 | 5.36 | 4.48 | 4.38 | 4.76 | 4.46 | 4.84 | 4.5 | 4.53 | 4.55 | 4.584 | 4.59 | 6.1 | 4.36 | 4.06 | 4.8 | [8 empty] | 4.5275
| 1 | 5 | 4.02 | 5.48 | 4.38 | 4.63 | 4.78 | 4.62 | 5.21 | 4.41 | 4.42 | 4.46 | 4.603 | 4.68 | 6.1 | 4.35 | 4.18 | 4.9 | [8 empty] | 85
| 1 | 6 | 4.16 | 5.05 | 4.5 | 4.53 | 4.55 | 4.56 | 4.78 | 4.5 | 4.53 | 4.42 | 4.5956 | 4.58 | 6.3 | 4.59 | 4.35 | 5
| Mean |  | 4.0517 | 5.1133 | 4.5017 | 4.4067 | 4.6983 | 4.5117 | 5.015 | 4.5017 | 4.5083 | 4.465 | 4.5975 | 4.57 | 6.2 | 4.525 | 4.2233 | 4.85
| Median |  | 4.06 | 5.055 | 4.5 | 4.4 | 4.735 | 4.52 | 5.005 | 4.5 | 4.525 | 4.45 | 4.5993 | 4.585 | 6.2 | 4.585 | 4.245 | 4.8
| Std Dev. |  | 0.095376 | 0.2575 | 0.086352 | 0.15706 | 0.090646 | 0.080602 | 0.19066 | 0.065549 | 0.079352 | 0.048477 | 0.015525 | 0.095917 | 0.089443 | 0.13457 | 0.13938 | 0.083666
| Rel.Std.Dev. |  | 2.354% | 5.036% | 1.918% | 3.564% | 1.929% | 1.787% | 3.802% | 1.456% | 1.76% | 1.086% | 0.3377% | 2.099% | 1.443% | 2.974% | 3.3% | 1.725%
| PDM3 |  | -10.51% | 12.94% | -0.5697% | -2.668% | 3.774% | -0.3489% | 10.77% | -0.5697% | -0.4225% | -1.38% | 1.548% | 0.9396% | 36.94% | -0.05435% | -6.717% | 7.124%
| Z-Score (Absolute) |  | 3.3 | 4.2 | 0.12 | 0.79 | 1.27 | 0.05 | 3.51 | 0.12 | 0.07 | 0.38 | 0.56 | 0.37 | 11.88 | 0.05 | 2.08 | 2.34
| Analytical results for K in OREAS 611 (Certified Value 0.201 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.2977 | 0.21 | 0.3 | 0.19 | 0.21 | 0.19 | 0.2 | 0.2 | 0.271 | 0.179 | 0.29095 | 0.13 | 0.2 | 0.17 | 0.2106 | 0.18 | 0.203 | 0.22 | 0.2 | 0.19 | 0.18 | 0.2 | 0.2585 |  | 1
| 1 | 2 | 0.3155 | 0.21 | 0.3 | 0.2 | 0.21 | 0.2 | 0.21 | 0.2 | 0.278 | 0.205 | 0.29282 | 0.13 | 0.19 | 0.16 | 0.208 | 0.18 | 0.206 | 0.2 | 0.21 | 0.2 | 0.18 | 0.2 | 0.2565 |  | #N/A
| 1 | 3 | 0.3342 | 0.21 | 0.3 | 0.19 | 0.21 | 0.2 | 0.2 | 0.2 | 0.263 | 0.206 | 0.29087 | 0.13 | 0.21 | 0.17 | 0.20478 | 0.17 | 0.208 | 0.21 | 0.22 | 0.2 | 0.18 | 0.2 | 0.2545 |  | 16
| 1 | 4 | 0.3201 | 0.21 | 0.28 | 0.19 | 0.22 | 0.19 | 0.21 | 0.2 | 0.271 | 0.212 | 0.28767 | 0.13 | 0.21 | 0.18 | 0.20284 | 0.18 | 0.205 | 0.2 | 0.21 | 0.2 | 0.18 | 0.21 | 0.2525 |  | 0.2007
| 1 | 5 | 0.3136 | 0.22 | 0.28 | 0.2 | 0.22 | 0.21 | 0.21 | 0.2 | 0.268 | 0.206 | 0.29657 | 0.13 | 0.21 | 0.17 | 0.21255 | 0.17 | 0.206 | 0.21 | 0.21 | 0.2 | 0.19 | 0.21 | 0.2505 |  | 86
| 1 | 6 | 0.3123 | 0.21 | 0.29 | 0.2 | 0.2 | 0.21 | 0.19 | 0.19 | 0.277 | 0.178 | 0.29407 | 0.12 | 0.2 | 0.17 | 0.19757 | 0.18 | 0.207 | 0.2 | 0.21 | 0.2 | 0.18 | 0.21 | 0.2485
| Mean |  | 0.31557 | 0.21167 | 0.29167 | 0.195 | 0.21167 | 0.2 | 0.20333 | 0.19833 | 0.27133 | 0.19767 | 0.29216 | 0.12833 | 0.20333 | 0.17 | 0.20606 | 0.17667 | 0.20583 | 0.20667 | 0.21 | 0.19833 | 0.18167 | 0.205 | 0.2535
| Median |  | 0.31455 | 0.21 | 0.295 | 0.195 | 0.21 | 0.2 | 0.205 | 0.2 | 0.271 | 0.2055 | 0.29188 | 0.13 | 0.205 | 0.17 | 0.20639 | 0.18 | 0.206 | 0.205 | 0.21 | 0.2 | 0.18 | 0.205 | 0.2535
| Std Dev. |  | 0.011845 | 0.0040825 | 0.0098319 | 0.0054772 | 0.0075277 | 0.0089443 | 0.008165 | 0.0040825 | 0.0056095 | 0.015055 | 0.0030616 | 0.0040825 | 0.008165 | 0.0063246 | 0.005488 | 0.005164 | 0.0017224 | 0.008165 | 0.0063246 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0037417
| Rel.Std.Dev. |  | 3.754% | 1.929% | 3.371% | 2.809% | 3.556% | 4.472% | 4.016% | 2.058% | 2.067% | 7.617% | 1.048% | 3.181% | 4.016% | 3.72% | 2.663% | 2.923% | 0.8368% | 3.951% | 3.012% | 2.058% | 2.247% | 2.672% | 1.476%
| PDM3 |  | 57.23% | 5.463% | 45.32% | -2.841% | 5.463% | -0.3495% | 1.311% | -1.18% | 35.19% | -1.512% | 45.57% | -36.06% | 1.311% | -15.3% | 2.668% | -11.98% | 2.557% | 2.972% | 4.633% | -1.18% | -9.484% | 2.142% | 26.31%
| Z-Score (Absolute) |  | 11.18 | 0.67 | 8.77 | 1.01 | 0.67 | 0.51 | 0.17 | 0.67 | 6.71 | 0.74 | 8.82 | 7.75 | 0.17 | 3.54 | 0.11 | 2.87 | 0.08 | 0.17 | 0.51 | 0.67 | 2.36 | 0 | 4.91
| Analytical results for La in OREAS 611 (Certified Value 6.86 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 2
|  |  | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  | 3
| 1 | 1 | 6.9 | 7.2 | 6.1 | 6.4 | 6.6 | 7.3 | 6.2 | 12.4 | 7.128 | 10.2704 | 5 | 6.1 | 5.8 | 6.0574 | 8.3 | 6.88 | 7.8 | 6.7 | 6.6 | 7 | 5 | 8.78 |  |  | 1
| 1 | 2 | 6.5 | 8 | 6.4 | 6.7 | 6.8 | 7.4 | 6.8 | 12.7 | 6.764 | 10.0448 | 5 | 6.5 | 6 | 5.8794 | 8.2 | 6.97 | 7.3 | 6.6 | 6.5 | 7 | 5 | 9.49 |  |  | 32
| 1 | 3 | 6.6 | 8.2 | 6.2 | 6.5 | 6.8 | 7.5 | 6.6 | 12.5 | 6.731 | 10.085 | 6 | 7.2 | 6 | 5.5937 | 8 | 7.09 | 7.8 | 7 | 6.7 | 7 | 5 | 9.2 |  |  | 16
| 1 | 4 | 6.9 | 8 | 6.4 | 6.8 | 6.7 | 7.3 | 6.2 | 12.5 | 6.845 | 10.0369 | 6 | 6.8 | 6.1 | 5.7585 | 8.3 | 7.34 | 7.2 | 7.1 | 6.9 | 7 | 5 | 9.91 |  |  | 6.8578
| 1 | 5 | 7.3 | 7.7 | 6.6 | 6.5 | 7.4 | 7.3 | 6.2 | 12.3 | 6.76 | 10.3359 | 5 | 6.9 | 5.8 | 5.9417 | 7.9 | 7.2 | 7.5 | 7.1 | 6.9 | 7 | 5 | 9.62 |  |  | 87
| 1 | 6 | 6.5 | 8 | 6.3 | 6.4 | 7.1 | 6.7 | 6.2 | 12.4 | 6.67 | 10.244 | 6 | 6.5 | 5.8 | 5.491 | 8.2 | 7.2 | 7.4 | 6.9 | 6.9 | 6 | 5 | 10.33
| Mean |  | 6.7833 | 7.85 | 6.3333 | 6.55 | 6.9 | 7.25 | 6.3667 | 12.4667 | 6.8163 | 10.1695 | 5.5 | 6.6667 | 5.9167 | 5.787 | 8.15 | 7.1133 | 7.5 | 6.9 | 6.75 | 6.8333 | 5 | 9.555
| Median |  | 6.75 | 8 | 6.35 | 6.5 | 6.8 | 7.3 | 6.2 | 12.45 | 6.762 | 10.1645 | 5.5 | 6.65 | 5.9 | 5.819 | 8.2 | 7.145 | 7.45 | 6.95 | 6.8 | 7 | 5 | 9.555
| Std Dev. |  | 0.31252 | 0.35637 | 0.17512 | 0.16432 | 0.29665 | 0.28107 | 0.26583 | 0.13663 | 0.16283 | 0.12939 | 0.54772 | 0.38297 | 0.13292 | 0.2151 | 0.16432 | 0.16848 | 0.25298 | 0.20976 | 0.17607 | 0.40825 | 0 | 0.54069
| Rel.Std.Dev. |  | 4.607% | 4.54% | 2.765% | 2.509% | 4.299% | 3.877% | 4.175% | 1.096% | 2.389% | 1.272% | 9.959% | 5.745% | 2.246% | 3.717% | 2.016% | 2.369% | 3.373% | 3.04% | 2.608% | 5.974% | 0% | 5.659%
| PDM3 |  | -1.085% | 14.47% | -7.647% | -4.488% | 0.6158% | 5.72% | -7.161% | 81.79% | -0.6042% | 48.29% | -19.8% | -2.787% | -13.72% | -15.61% | 18.84% | 3.727% | 9.365% | 0.6158% | -1.571% | -0.3563% | -27.09% | 39.33%
| Z-Score (Absolute) |  | 0.15 | 1.2 | 0.72 | 0.44 | 0 | 0.44 | 0.67 | 7.04 | 0.11 | 4.13 | NA | 0.3 | 1.24 | 1.41 | 1.58 | 0.27 | 0.76 | 0 | 0.19 | NA | NA | 3.36
Results from laboratories 13, 23 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 611 (Certified Value 8.25 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g |  |  |  |  | 3
| 1 | 1 | 7.4 | 9.8 | 10.3 | 8 | 8.2 | 8.4 | 7.5 | 8.3 | 10.9 | 8.24 | 10.5066 | 7.9 | 8.927 | 7.6 | 8.487 | 9 | 8 | 8 | 8 | 9 |  |  |  |  | 1
| 1 | 2 | 7.8 | 9.7 | 10.6 | 8.4 | 8.4 | 8.2 | 7.4 | 8.5 | 10.9 | 8.12 | 10.2692 | 7.7 | 8.8341 | 7.7 | 7.543 | 9 | 8 | 8 | 9.1 | 9 |  |  |  |  | #N/A
| 1 | 3 | 8 | 9.5 | 10.7 | 8.3 | 8.6 | 8.2 | 7.1 | 8.4 | 10.4 | 8.22 | 10.7206 | 7.8 | 8.8573 | 7.4 | 7.709 | 8 | 8 | 8 | 8.2 | 9 |  |  |  |  | 16
| 1 | 4 | 7.9 | 9.7 | 10.4 | 8.2 | 8.6 | 8.3 | 6.7 | 8.3 | 10.7 | 8.55 | 10.2481 | 8.2 | 8.9058 | 7.9 | 7.916 | 9 | 8 | 8 | 8.3 | 9 |  |  |  |  | 8.2534
| 1 | 5 | 7.8 | 10.6 | 10.1 | 8.2 | 8.4 | 8.7 | 6.8 | 8.2 | 10.3 | 8.28 | 10.7764 | 8 | 8.9456 | 7.3 | 7.727 | 8 | 8 | 8 | 8.3 | 9 |  |  |  |  | 88
| 1 | 6 | 7.9 | 9.7 | 10.2 | 8.3 | 8.3 | 8.4 | 8.2 | 8.4 | 10.6 | 8.24 | 10.6048 | 7.7 | 8.5163 | 7.5 | 8.385 | 9 | 9 | 8 | 9.5 | 9
| Mean |  | 7.8 | 9.8333 | 10.3833 | 8.2333 | 8.4167 | 8.3667 | 7.2833 | 8.35 | 10.6333 | 8.275 | 10.5209 | 7.8833 | 8.831 | 7.5667 | 7.9612 | 8.6667 | 8.1667 | 8 | 8.5667 | 9
| Median |  | 7.85 | 9.7 | 10.35 | 8.25 | 8.4 | 8.35 | 7.25 | 8.35 | 10.65 | 8.24 | 10.5557 | 7.85 | 8.8816 | 7.55 | 7.8215 | 9 | 8 | 8 | 8.3 | 9
| Std Dev. |  | 0.20976 | 0.38816 | 0.23166 | 0.13663 | 0.16021 | 0.18619 | 0.54924 | 0.10488 | 0.25033 | 0.14502 | 0.22366 | 0.19408 | 0.1598 | 0.21602 | 0.38768 | 0.5164 | 0.40825 | 0 | 0.59217 | 0
| Rel.Std.Dev. |  | 2.689% | 3.947% | 2.231% | 1.659% | 1.903% | 2.225% | 7.541% | 1.256% | 2.354% | 1.752% | 2.126% | 2.462% | 1.81% | 2.855% | 4.87% | 5.958% | 4.999% | 0% | 6.913% | 0%
| PDM3 |  | -5.493% | 19.14% | 25.81% | -0.2428% | 1.979% | 1.373% | -11.75% | 1.171% | 28.84% | 0.2621% | 27.47% | -4.483% | 6.999% | -8.32% | -3.54% | 5.008% | -1.051% | -3.07% | 3.796% | 9.046%
| Z-Score (Absolute) |  | 0.79 | 2.1 | 2.88 | 0.18 | 0.08 | 0.01 | 1.53 | 0.01 | 3.24 | 0.12 | 3.08 | 0.68 | 0.67 | 1.13 | 0.57 | NA | NA | NA | 0.3 | NA
Results from laboratories 20, 21, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 611 (Indicative Value 31.7 ppb) | [25 empty] | Indicative
Lu | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab |  |  |  |  | 1
| No | No | 2.02: 2.04; 2.03: 2.20; 2.04: 2.21 |  |  |  |  | ppb
|  |  | 2.02: AR*MS; 2.03: AR*OES/MS; 2.04: AR*OES/MS |  |  |  |  | 1
|  |  | 2.02: 0.5g; 2.03: 0.25g; 2.04: 0.25g |  |  |  |  | 1
| 1 | 1 | 2.02: < 100; 2.03: 40; 2.04: 30 |  |  |  |  | 1
| 1 | 2 | 2.02: < 100; 2.03: 30; 2.04: 30 |  |  |  |  | 6
| 1 | 3 | 2.02: < 100; 2.03: 30; 2.04: 30 |  |  |  |  | 16
| 1 | 4 | 2.02: < 100; 2.03: 40; 2.04: 30 |  |  |  |  | 31.6667
| 1 | 5 | 2.02: < 100; 2.03: 30; 2.04: 30 |  |  |  |  | 15
| 1 | 6 | 2.02: < 100; 2.03: 30; 2.04: 30
| Mean |  | 2.03: 33.3333; 2.04: 30
| Median |  | 2.03: 30; 2.04: 30
| Std Dev. |  | 2.03: 5.164; 2.04: 0
| Rel.Std.Dev. |  | 2.03: 15.49%; 2.04: 0%
| PDM3 |  | 2.03: 5.263%; 2.04: -5.263%
| Z-Score (Absolute) |  | 2.02: 3.37; 2.03: 0; 2.04: 0.67
| Analytical results for Mg in OREAS 611 (Certified Value 1000 ppm) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 1100 | 900 | 1100 | 900 | 1000 | 900 | 1000 | 900 | 1200 | 880 | 1129.5835 | 800 | 1000 | 1000 | 1005.5962 | 1000 | 1000 | 1100 | 1100 | 900 | 1000 | 1000 | 1128 |  | 1
| 1 | 2 | 1100 | 1000 | 1100 | 900 | 1000 | 900 | 1000 | 1000 | 1200 | 990 | 1127.5236 | 800 | 1000 | 1000 | 996.8009 | 1000 | 980 | 1000 | 1000 | 1000 | 1000 | 1000 | 1173 |  | #N/A
| 1 | 3 | 1100 | 1000 | 1100 | 900 | 1000 | 900 | 1000 | 1000 | 1200 | 990 | 1126.2843 | 800 | 1000 | 1000 | 997.5564 | 1000 | 990 | 1200 | 1100 | 1000 | 1000 | 1000 | 1218 |  | 16
| 1 | 4 | 1100 | 1000 | 1000 | 900 | 1000 | 900 | 1000 | 900 | 1200 | 1030 | 1089.7645 | 800 | 1000 | 1000 | 1002.8196 | 1000 | 970 | 1000 | 1000 | 900 | 1000 | 1000 | 1263 |  | 999.8813
| 1 | 5 | 1100 | 1000 | 1000 | 900 | 1000 | 900 | 1000 | 1000 | 1200 | 1000 | 1080.2339 | 800 | 1000 | 1000 | 1004.5028 | 1000 | 1000 | 1100 | 1100 | 900 | 1000 | 900 | 1208 |  | 89
| 1 | 6 | 1100 | 900 | 1100 | 1000 | 900 | 900 | 900 | 900 | 1200 | 860 | 1132.6565 | 800 | 1000 | 1000 | 972.439 | 1000 | 1030 | 1000 | 1100 | 1000 | 1000 | 1000 | 1253
| Mean |  | 1100 | 966.6667 | 1066.6667 | 916.6667 | 983.3333 | 900 | 983.3333 | 950 | 1200 | 958.3333 | 1114.341 | 800 | 1000 | 1000 | 996.6192 | 1000 | 995 | 1066.6667 | 1066.6667 | 950 | 1000 | 983.3333 | 1207.1667
| Median |  | 1100 | 1000 | 1100 | 900 | 1000 | 900 | 1000 | 950 | 1200 | 990 | 1126.9039 | 800 | 1000 | 1000 | 1000.188 | 1000 | 995 | 1050 | 1100 | 950 | 1000 | 1000 | 1213
| Std Dev. |  | 0 | 51.6398 | 51.6398 | 40.8248 | 40.8248 | 0 | 40.8248 | 54.7723 | 0 | 70.2614 | 23.0283 | 0 | 0 | 0 | 12.3837 | 0 | 20.7364 | 81.6497 | 51.6398 | 54.7723 | 0 | 40.8248 | 50.5388
| Rel.Std.Dev. |  | 0% | 5.342% | 4.841% | 4.454% | 4.152% | 0% | 4.152% | 5.766% | 0% | 7.332% | 2.067% | 0% | 0% | 0% | 1.243% | 0% | 2.084% | 7.655% | 4.841% | 5.766% | 0% | 4.152% | 4.187%
| PDM3 |  | 10.01% | -3.322% | 6.679% | -8.322% | -1.655% | -9.989% | -1.655% | -4.989% | 20.01% | -4.155% | 11.45% | -19.99% | 0.01187% | 0.01187% | -0.3263% | 0.01187% | -0.4882% | 6.679% | 6.679% | -4.989% | 0.01187% | -1.655% | 20.73%
| Z-Score (Absolute) |  | 1.5 | 0.43 | 1.01 | 1.16 | 0.19 | 1.4 | 0.19 | 0.67 | 2.94 | 0.55 | 1.7 | 2.84 | 0.05 | 0.05 | 0 | 0.05 | 0.02 | 1.01 | 1.01 | 0.67 | 0.05 | 0.19 | 3.05
| Analytical results for Mn in OREAS 611 (Certified Value 62 ppm) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 1
| 1 | 1 | 48 | 67 | 66 | 62 | 57 | 61 | 61 | 62 | 61 | 66 | 55.8 | 55.8942 | 51 | 58 | 64 | 61.4967 | 59 | 61.08 | 64 | 68 | 59 | 58 | 87 | 75.22 | 1
| 1 | 2 | 42.7 | 70 | 68 | 65 | 58 | 63 | 60 | 62 | 63 | 68 | 58 | 63.376 | 50 | 58 | 62 | 61.4918 | 60 | 60.607 | 61 | 64 | 60 | 59 | 88 | 82.43 | 34
| 1 | 3 | 44.8 | 69 | 67 | 64 | 58 | 65 | 60 | 61 | 62 | 65 | 58.1 | 62.2792 | 50 | 60 | 63 | 62.1496 | 58 | 60.936 | 60 | 68 | 60 | 59 | 87 | 89.64 | 16
| 1 | 4 | 44.6 | 70 | 67 | 61 | 58 | 65 | 60 | 64 | 62 | 67 | 59.7 | 60.9486 | 50 | 61 | 66 | 61.6536 | 59 | 60.132 | 59 | 65 | 59 | 59 | 90 | 90.85 | 61.9307
| 1 | 5 | 50.4 | 70 | 74 | 73 | 58 | 64 | 62 | 63 | 62 | 65 | 56.8 | 56.7139 | 50 | 60 | 66 | 61.5544 | 58 | 60.73 | 58 | 67 | 59 | 60 | 87 | 81.06 | 90
| 1 | 6 | 45.1 | 70 | 69 | 63 | 59 | 64 | 61 | 60 | 61 | 67 | 53.2 | 56.4581 | 50 | 59 | 63 | 59.3871 | 59 | 62.8 | 61 | 68 | 59 | 60 | 87 | 78.27
| Mean |  | 45.9333 | 69.3333 | 68.5 | 64.6667 | 58 | 63.6667 | 60.6667 | 62 | 61.8333 | 66.3333 | 56.9333 | 59.2783 | 50.1667 | 59.3333 | 64 | 61.2889 | 58.8333 | 61.0475 | 60.5 | 66.6667 | 59.3333 | 59.1667 | 87.6667 | 82.9117
| Median |  | 44.95 | 70 | 67.5 | 63.5 | 58 | 64 | 60.5 | 62 | 62 | 66.5 | 57.4 | 58.8312 | 50 | 59.5 | 63.5 | 61.5255 | 59 | 60.833 | 60.5 | 67.5 | 59 | 59 | 87 | 81.745
| Std Dev. |  | 2.7725 | 1.2111 | 2.881 | 4.3205 | 0.63246 | 1.5055 | 0.8165 | 1.4142 | 0.75277 | 1.2111 | 2.2536 | 3.3036 | 0.40825 | 1.2111 | 1.6733 | 0.96389 | 0.75277 | 0.91845 | 2.0736 | 1.7512 | 0.5164 | 0.75277 | 1.2111 | 6.2073
| Rel.Std.Dev. |  | 6.036% | 1.747% | 4.206% | 6.681% | 1.09% | 2.365% | 1.346% | 2.281% | 1.217% | 1.826% | 3.958% | 5.573% | 0.8138% | 2.041% | 2.615% | 1.573% | 1.28% | 1.504% | 3.428% | 2.627% | 0.8703% | 1.272% | 1.381% | 7.487%
| PDM3 |  | -25.83% | 11.95% | 10.61% | 4.418% | -6.347% | 2.803% | -2.041% | 0.1118% | -0.1573% | 7.109% | -8.069% | -4.283% | -19% | -4.194% | 3.341% | -1.036% | -5.001% | -1.426% | -2.31% | 7.647% | -4.194% | -4.463% | 41.56% | 33.88%
| Z-Score (Absolute) |  | 3.85 | 2.07 | 1.86 | 0.89 | 0.8 | 0.63 | 0.13 | 0.21 | 0.17 | 1.31 | 1.07 | 0.48 | 2.78 | 0.46 | 0.72 | 0.03 | 0.59 | 0.03 | 0.17 | 1.39 | 0.46 | 0.51 | 6.7 | 5.5
| Analytical results for Mo in OREAS 611 (Certified Value 5.13 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 4.9 | 5.4 | 5.27 | 4.84 | 4.93 | 5.12 | 5.31 | 5.12 | 5 | 5.4 | 5.3932 | 5 | 4.27 | 5.96 | 5.7235 | 4.67 | 5.177 | 5.62 | 5.1 | 5 | 5 | 5 | 9.675 |  | 1
| 1 | 2 | 4.9 | 5.66 | 5.46 | 4.69 | 5.15 | 4.98 | 5.38 | 5.29 | 5 | 5.27 | 5.3855 | 5 | 4.48 | 5.85 | 6.2389 | 4.9 | 5.013 | 5.96 | 4.61 | 5 | 5.1 | 5 | 9.687 |  | 35
| 1 | 3 | 5 | 5.59 | 5.13 | 4.65 | 5.72 | 5.04 | 5.2 | 5.1 | 5 | 5.27 | 5.2583 | 5 | 4.89 | 5.68 | 5.2469 | 4.46 | 5.895 | 5.6 | 4.97 | 5 | 5.4 | 5 | 9.593 |  | 16
| 1 | 4 | 5 | 5.56 | 5.02 | 5.07 | 5.11 | 5.07 | 5.23 | 4.94 | 5 | 5.3 | 5.1214 | 5 | 4.94 | 6.31 | 6.0949 | 4.71 | 5.181 | 5.53 | 4.65 | 5 | 5 | 5 | 9.595 |  | 5.1262
| 1 | 5 | 4.9 | 5.79 | 5.13 | 5.14 | 5.35 | 5.24 | 5.08 | 4.94 | 4.8 | 5.31 | 5.1421 | 5 | 4.67 | 6.06 | 5.796 | 4.73 | 5.566 | 5.59 | 4.85 | 5 | 5 | 5 | 9.678 |  | 91
| 1 | 6 | 5 | 5.54 | 5.04 | 5.62 | 4.86 | 5.05 | 5.02 | 5.47 | 4.8 | 5.27 | 5.0395 | 4 | 4.51 | 6.17 | 5.1177 | 4.43 | 5.498 | 5.83 | 4.89 | 5 | 5.2 | 5 | 9.69
| Mean |  | 4.95 | 5.59 | 5.175 | 5.0017 | 5.1867 | 5.0833 | 5.2033 | 5.1433 | 4.9333 | 5.3033 | 5.2233 | 4.8333 | 4.6267 | 6.005 | 5.703 | 4.65 | 5.3883 | 5.6883 | 4.845 | 5 | 5.1167 | 5 | 9.653
| Median |  | 4.95 | 5.575 | 5.13 | 4.955 | 5.13 | 5.06 | 5.215 | 5.11 | 5 | 5.285 | 5.2002 | 5 | 4.59 | 6.01 | 5.7598 | 4.69 | 5.3395 | 5.61 | 4.87 | 5 | 5.05 | 5 | 9.6765
| Std Dev. |  | 0.054772 | 0.12992 | 0.16526 | 0.36141 | 0.31335 | 0.089144 | 0.13574 | 0.20656 | 0.10328 | 0.050465 | 0.14639 | 0.40825 | 0.25758 | 0.22581 | 0.44732 | 0.17743 | 0.32522 | 0.16798 | 0.18759 | 0 | 0.16021 | 0 | 0.046039
| Rel.Std.Dev. |  | 1.107% | 2.324% | 3.193% | 7.226% | 6.041% | 1.754% | 2.609% | 4.016% | 2.094% | 0.9516% | 2.803% | 8.447% | 5.567% | 3.76% | 7.844% | 3.816% | 6.036% | 2.953% | 3.872% | 0% | 3.131% | 0% | 0.4769%
| PDM3 |  | -3.437% | 9.048% | 0.9523% | -2.429% | 1.18% | -0.8359% | 1.505% | 0.3346% | -3.762% | 3.456% | 1.895% | -5.713% | -9.744% | 17.14% | 11.25% | -9.289% | 5.114% | 10.97% | -5.485% | -2.461% | -0.1856% | -2.461% | 88.31%
| Z-Score (Absolute) |  | 0.67 | 1.56 | 0.11 | 0.49 | 0.15 | 0.21 | 0.21 | 0 | 0.73 | 0.56 | 0.28 | 1.08 | 1.8 | 3.01 | 1.95 | 1.72 | 0.85 | 1.9 | 1.04 | 0.5 | 0.09 | 0.5 | 15.73
Results from laboratories 13, 22 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 611 (Certified Value 0.044 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.05 | 0.044 | 0.068 | 0.04 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.038 | 0.039758 | 0.04 | 0.043 | 0.03 | 0.059724 | 0.03 | 0.034 | 0.04 | 0.04 | 0.04 | 0.08 | 0.05 | 0.06588 |  | 1
| 1 | 2 | 0.06 | 0.043 | 0.071 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.06 | 0.044 | 0.041488 | 0.04 | 0.041 | 0.03 | 0.059112 | 0.03 | 0.033 | 0.05 | 0.04 | 0.05 | 0.07 | 0.05 | 0.07305 |  | #N/A
| 1 | 3 | 0.06 | 0.044 | 0.075 | 0.05 | 0.05 | 0.04 | 0.04 | 0.05 | 0.05 | 0.044 | 0.041775 | 0.04 | 0.045 | 0.03 | 0.057571 | 0.03 | 0.032 | 0.06 | 0.05 | 0.05 | 0.08 | 0.05 | 0.08022 |  | 16
| 1 | 4 | 0.06 | 0.045 | 0.068 | 0.05 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.045 | 0.039119 | 0.04 | 0.045 | 0.04 | 0.056192 | 0.03 | 0.033 | 0.05 | 0.04 | 0.05 | 0.08 | 0.05 | 0.08439 |  | 0.044336
| 1 | 5 | 0.06 | 0.045 | 0.068 | 0.05 | 0.05 | 0.04 | 0.05 | 0.04 | 0.05 | 0.044 | 0.041555 | 0.04 | 0.044 | 0.03 | 0.060043 | 0.03 | 0.033 | 0.05 | 0.04 | 0.05 | 0.08 | 0.05 | 0.09446 |  | 92
| 1 | 6 | 0.06 | 0.044 | 0.068 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.06 | 0.039 | 0.040578 | 0.04 | 0.044 | 0.03 | 0.054441 | 0.03 | 0.033 | 0.04 | 0.04 | 0.05 | 0.06 | 0.05 | 0.070173
| Mean |  | 0.058333 | 0.044167 | 0.069667 | 0.048333 | 0.05 | 0.04 | 0.043333 | 0.041667 | 0.053333 | 0.042333 | 0.040712 | 0.04 | 0.043667 | 0.031667 | 0.057847 | 0.03 | 0.033 | 0.048333 | 0.041667 | 0.048333 | 0.075 | 0.05 | 0.078029
| Median |  | 0.06 | 0.044 | 0.068 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.044 | 0.041033 | 0.04 | 0.044 | 0.03 | 0.058342 | 0.03 | 0.033 | 0.05 | 0.04 | 0.05 | 0.08 | 0.05 | 0.076635
| Std Dev. |  | 0.0040825 | 0.00075277 | 0.0028752 | 0.0040825 | 7.6012e-18 | 0 | 0.005164 | 0.0040825 | 0.005164 | 0.0030111 | 0.001087 | 0 | 0.0015055 | 0.0040825 | 0.0022073 | 0 | 0.00063246 | 0.0075277 | 0.0040825 | 0.0040825 | 0.0083666 | 7.6012e-18 | 0.010482
| Rel.Std.Dev. |  | 6.999% | 1.704% | 4.127% | 8.447% | 1.52e-14% | 0% | 11.92% | 9.798% | 9.682% | 7.113% | 2.67% | 0% | 3.448% | 12.89% | 3.816% | 0% | 1.917% | 15.57% | 9.798% | 8.447% | 11.16% | 1.52e-14% | 13.43%
| PDM3 |  | 31.57% | -0.3826% | 57.13% | 9.015% | 12.77% | -9.78% | -2.262% | -6.021% | 20.29% | -4.518% | -8.174% | -9.78% | -1.51% | -28.58% | 30.47% | -32.34% | -25.57% | 9.015% | -6.021% | 9.015% | 69.16% | 12.77% | 75.99%
| Z-Score (Absolute) |  | 2.29 | 0 | 4.13 | 0.67 | 0.94 | 0.67 | 0.13 | 0.4 | 1.48 | 0.3 | 0.56 | 0.67 | 0.08 | 2.02 | 2.21 | 2.29 | 1.81 | 0.67 | 0.4 | 0.67 | 4.99 | 0.94 | 5.48
| Analytical results for Nb in OREAS 611 (Certified Value 0.17 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [7 empty] | 2
| 1 | 1 | < 0.2 | 0.07 | 0.3 | 0.15 | 0.16 | 0.14 | 0.17 | 0.16 | 0.4 | 0.12 | < 0.2 | 0.33 | < 1 | < 0.2 | 0.22 | 0.13 | 0.1 | [7 empty] | 1
| 1 | 2 | < 0.2 | 0.07 | 0.3 | 0.17 | 0.19 | 0.16 | 0.17 | 0.16 | 0.4 | 0.12 | < 0.2 | 0.34 | < 1 | < 0.2 | 0.21 | 0.23 | 0.1 | [7 empty] | 36
| 1 | 3 | 0.2 | 0.07 | 0.3 | 0.17 | 0.19 | 0.15 | 0.14 | 0.16 | 0.4 | 0.13 | < 0.2 | 0.31 | < 1 | < 0.2 | 0.22 | 0.18 | 0.1 | [7 empty] | 16
| 1 | 4 | < 0.2 | 0.07 | 0.3 | 0.18 | 0.2 | 0.15 | 0.18 | 0.15 | 0.4 | 0.13 | < 0.2 | 0.32 | < 1 | < 0.2 | 0.22 | 0.19 | 0.2 | [7 empty] | 0.16821
| 1 | 5 | < 0.2 | 0.09 | 0.3 | 0.18 | 0.17 | 0.16 | 0.18 | 0.15 | 0.4 | 0.13 | < 0.2 | 0.32 | < 1 | < 0.2 | 0.23 | 0.15 | 0.1 | [7 empty] | 93
| 1 | 6 | < 0.2 | 0.07 | 0.3 | 0.14 | 0.19 | 0.15 | 0.17 | 0.16 | 0.4 | 0.12 | < 0.2 | 0.32 | < 1 | < 0.2 | 0.22 | 0.14 | 0.2
| Mean |  | 0.2 | 0.073333 | 0.3 | 0.165 | 0.18333 | 0.15167 | 0.16833 | 0.15667 | 0.4 | 0.125 |  | 0.32333 |  |  | 0.22 | 0.17 | 0.13333
| Median |  | 0.2 | 0.07 | 0.3 | 0.17 | 0.19 | 0.15 | 0.17 | 0.16 | 0.4 | 0.125 |  | 0.32 |  |  | 0.22 | 0.165 | 0.1
| Std Dev. |  |  | 0.008165 | 0 | 0.016432 | 0.015055 | 0.0075277 | 0.01472 | 0.005164 | 6.0809e-17 | 0.0054772 |  | 0.010328 |  |  | 0.0063246 | 0.037417 | 0.05164
| Rel.Std.Dev. |  |  | 11.13% | 0% | 9.959% | 8.212% | 4.963% | 8.744% | 3.296% | 1.52e-14% | 4.382% |  | 3.194% |  |  | 2.875% | 22.01% | 38.73%
| PDM3 |  | 18.9% | -56.4% | 78.35% | -1.907% | 8.992% | -9.834% | 0.07431% | -6.862% | 137.8% | -25.69% |  | 92.22% |  |  | 30.79% | 1.065% | -20.73%
| Z-Score (Absolute) |  | NA | 1.55 | NA | 0 | 0.31 | 0.22 | 0.06 | 0.14 | NA | 0.67 | 1.1 | 2.67 | 5.65 | 1.1 | 0.93 | 0.08 | NA
| Analytical results for Nd in OREAS 611 (Indicative Value 8.33 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | 2.02: Lab | [7 empty] | 1
| No | No | 2.02: 2.04 | [7 empty] | ppm
|  |  | 2.02: AR*MS | [7 empty] | 2
|  |  | 2.02: 0.5g | [7 empty] | 2
| 1 | 1 | 2.02: 7.85 | [7 empty] | 1
| 1 | 2 | 2.02: 8.49 | [7 empty] | 10
| 1 | 3 | 2.02: 8.76 | [7 empty] | 16
| 1 | 4 | 2.02: 8.34 | [7 empty] | 8.325
| 1 | 5 | 2.02: 7.99 | [7 empty] | 16
| 1 | 6 | 2.02: 8.52
| Mean |  | 2.02: 8.325
| Median |  | 2.02: 8.415
| Std Dev. |  | 2.02: 0.34425
| Rel.Std.Dev. |  | 2.02: 4.135%
| PDM3 |  | 2.02: 2.22e-14%
Z-Score (Absolute) | bar: 2.02
| Analytical results for Ni in OREAS 611 (Certified Value 26.9 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 1
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 2
| 1 | 1 | 16.1 | 28 | 29.9 | 25 | 25 | 26.6 | 26.7 | 26.2 | 26.2 | 29 | 25.4 | 26.4858 | 29 | 23.3 | 27.6 | 24.1317 | 26.3 | 26.483 | 25.9 | 30.4 | 25 | 34.6 | 28 | 29.43 | 1
| 1 | 2 | 14 | 29 | 29.8 | 26.7 | 25 | 26.8 | 26 | 27.2 | 27.3 | 29 | 25.5 | 26.7578 | 29 | 23 | 27.5 | 24.1181 | 26.2 | 27.07 | 26 | 29 | 25 | 32.1 | 28 | 31.04 | 21
| 1 | 3 | 14 | 28 | 30.2 | 27.3 | 25.5 | 26.7 | 26.2 | 26.9 | 26.5 | 28 | 25.5 | 26.133 | 29 | 24.7 | 28 | 24.5744 | 25.4 | 27.355 | 25.9 | 30.2 | 26 | 31.1 | 28 | 32.65 | 16
| 1 | 4 | 15.2 | 29 | 30.3 | 26.2 | 24.7 | 27.3 | 26.1 | 27.1 | 26.1 | 29 | 26.4 | 26.0001 | 30 | 26.1 | 29 | 25.0119 | 26.8 | 26.886 | 24.2 | 29.6 | 26 | 34 | 28 | 34.26 | 26.9125
| 1 | 5 | 15.9 | 28 | 30.7 | 25.9 | 26.1 | 26.3 | 28 | 25.7 | 25.7 | 28 | 25.1 | 26.6773 | 30 | 24.4 | 27.2 | 25.2502 | 25.8 | 26.828 | 24.6 | 30.3 | 25 | 31.9 | 28 | 35.87 | 94
| 1 | 6 | 15.2 | 28 | 29.4 | 26.4 | 25.6 | 26.1 | 27.2 | 25.6 | 27.3 | 29 | 24.4 | 26.9635 | 29 | 24.3 | 28.8 | 23.031 | 26.8 | 27.913 | 26.2 | 31 | 26 | 34.2 | 28 | 30.48
| Mean |  | 15.0667 | 28.3333 | 30.05 | 26.25 | 25.3167 | 26.6333 | 26.7 | 26.45 | 26.5167 | 28.6667 | 25.3833 | 26.5029 | 29.3333 | 24.3 | 28.0167 | 24.3529 | 26.2167 | 27.0892 | 25.4667 | 30.0833 | 25.5 | 32.9833 | 28 | 32.2883
| Median |  | 15.2 | 28 | 30.05 | 26.3 | 25.25 | 26.65 | 26.45 | 26.55 | 26.35 | 29 | 25.45 | 26.5816 | 29 | 24.35 | 27.8 | 24.3531 | 26.25 | 26.978 | 25.9 | 30.25 | 25.5 | 33.05 | 28 | 31.845
| Std Dev. |  | 0.90259 | 0.5164 | 0.45056 | 0.77653 | 0.51153 | 0.41793 | 0.77974 | 0.71204 | 0.65853 | 0.5164 | 0.64936 | 0.37348 | 0.5164 | 1.1045 | 0.73326 | 0.79247 | 0.55287 | 0.49531 | 0.84301 | 0.69402 | 0.54772 | 1.458 | 0 | 2.4432
| Rel.Std.Dev. |  | 5.991% | 1.823% | 1.499% | 2.958% | 2.021% | 1.569% | 2.92% | 2.692% | 2.483% | 1.801% | 2.558% | 1.409% | 1.76% | 4.545% | 2.617% | 3.254% | 2.109% | 1.828% | 3.31% | 2.307% | 2.148% | 4.42% | 0% | 7.567%
| PDM3 |  | -44.02% | 5.28% | 11.66% | -2.462% | -5.93% | -1.037% | -0.7894% | -1.718% | -1.471% | 6.518% | -5.682% | -1.522% | 8.995% | -9.707% | 4.103% | -9.511% | -2.585% | 0.6566% | -5.372% | 11.78% | -5.248% | 22.56% | 4.041% | 19.98%
| Z-Score (Absolute) |  | 5.78 | 0.88 | 1.75 | 0.16 | 0.63 | 0.03 | 0.06 | 0.06 | 0.03 | 1.05 | 0.6 | 0.04 | 1.39 | 1.14 | 0.72 | 1.12 | 0.18 | 0.26 | 0.56 | 1.76 | 0.54 | 3.22 | 0.72 | 2.87
| Analytical results for P in OREAS 611 (Certified Value 243 ppm) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 245 | 250 | 300 | 220 | 240 | 230 | 230 | 230 | 273 | 199 | 607.3274 | 130 | 230 | 250 | 237.227 | 400 | 420 | 254 | 300 | 300 | 300 | 360 | 250.3 |  | 1
| 1 | 2 | 264 | 240 | 300 | 230 | 240 | 240 | 230 | 240 | 266 | 223 | 612.0631 | 110 | 230 | 250 | 233.7899 | 300 | 460 | 249 | 300 | 300 | 300 | 360 | 245.5 |  | #N/A
| 1 | 3 | 257 | 240 | 310 | 230 | 240 | 230 | 230 | 240 | 253 | 231 | 600.319 | 100 | 240 | 260 | 237.4574 | 300 | 440 | 253 | 300 | 300 | 300 | 350 | 250.07 |  | 16
| 1 | 4 | 266 | 240 | 300 | 230 | 250 | 230 | 230 | 240 | 268 | 231 | 598.9676 | 90 | 240 | 260 | 234.5986 | 400 | 480 | 255 | 300 | 300 | 300 | 360 | 235.9 |  | 242.618
| 1 | 5 | 254 | 260 | 300 | 230 | 260 | 240 | 240 | 240 | 260 | 224 | 613.4754 | 110 | 240 | 250 | 236.8513 | 400 | 450 | 254 | 300 | 300 | 300 | 350 | 251.1 |  | 95
| 1 | 6 | 257 | 240 | 290 | 240 | 240 | 240 | 220 | 240 | 270 | 194 | 619.8079 | 100 | 240 | 250 | 228.2689 | 400 | 510 | 249 | 300 | 300 | 290 | 350 | 240.63
| Mean |  | 257.1667 | 245 | 300 | 230 | 245 | 235 | 230 | 238.3333 | 265 | 217 | 608.6601 | 106.6667 | 236.6667 | 253.3333 | 234.6989 | 366.6667 | 460 | 252.3333 | 300 | 300 | 298.3333 | 355 | 245.5833
| Median |  | 257 | 240 | 300 | 230 | 240 | 235 | 230 | 240 | 267 | 223.5 | 609.6952 | 105 | 240 | 250 | 235.7249 | 400 | 455 | 253.5 | 300 | 300 | 300 | 355 | 247.785
| Std Dev. |  | 7.5211 | 8.3666 | 6.3246 | 6.3246 | 8.3666 | 5.4772 | 6.3246 | 4.0825 | 7.3212 | 16.3095 | 8.054 | 13.6626 | 5.164 | 5.164 | 3.4875 | 51.6398 | 31.6228 | 2.6583 | 0 | 0 | 4.0825 | 5.4772 | 6.1826
| Rel.Std.Dev. |  | 2.925% | 3.415% | 2.108% | 2.75% | 3.415% | 2.331% | 2.75% | 1.713% | 2.763% | 7.516% | 1.323% | 12.81% | 2.182% | 2.038% | 1.486% | 14.08% | 6.875% | 1.053% | 0% | 0% | 1.368% | 1.543% | 2.517%
| PDM3 |  | 5.997% | 0.9818% | 23.65% | -5.201% | 0.9818% | -3.14% | -5.201% | -1.766% | 9.225% | -10.56% | 150.9% | -56.04% | -2.453% | 4.417% | -3.264% | 51.13% | 89.6% | 4.004% | 23.65% | 23.65% | 22.96% | 46.32% | 1.222%
| Z-Score (Absolute) |  | 0.59 | 0.01 | 2.72 | 0.76 | 0.01 | 0.51 | 0.76 | 0.35 | 0.98 | 1.4 | 18.04 | 6.88 | 0.43 | 0.4 | 0.53 | NA | 10.66 | 0.35 | NA | NA | 2.63 | 5.45 | 0.01
Results from laboratories 18, 21 & 22 were removed due to their 0.01 wt.% reading resolution.
| Analytical results for Pb in OREAS 611 (Certified Value 481 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 0
| 1 | 1 | 412 | 485 | 490.46 | 443 | 455 | 479 | 454 | 481 | 473 | 569 | 495.7 | 503.3119 | 491 | 473.27 | 463 | 454.9906 | 481.2 | 465.6 | 497 | 438 | 433 | 492.8 | 490 | 530.4 | 1
| 1 | 2 | 396 | 504 | 494.55 | 462 | 457 | 480 | 452 | 479 | 483 | 565 | 478.6 | 504.986 | 507 | 470.2 | 469 | 453.2304 | 480.3 | 483.598 | 486.8 | 443 | 426 | 487.1 | 490 | 550.3 | 23
| 1 | 3 | 391 | 514 | 496.52 | 464 | 457 | 486 | 454 | 476 | 478 | 569 | 476.5 | 503.4951 | 505 | 491.18 | 481 | 456.4757 | 478.5 | 492.242 | 492.9 | 443 | 456 | 492.2 | 500 | 540.2 | 16
| 1 | 4 | 388 | 506 | 501.51 | 449 | 456 | 490 | 453 | 487 | 479 | 576 | 486.4 | 498.9983 | 517 | 495.31 | 483 | 454.4113 | 482.9 | 486.63 | 482.5 | 445 | 445 | 497.8 | 500 | 550.1 | 480.6527
| 1 | 5 | 414 | 498 | 524.87 | 449 | 455 | 494 | 463 | 484 | 481 | 574 | 475 | 508.1461 | 504 | 485.63 | 466 | 452.3838 | 469.7 | 498.642 | 481.5 | 450 | 444 | 495.8 | 490 | 550 | 96
| 1 | 6 | 398 | 504 | 503.47 | 446 | 464 | 464 | 459 | 464 | 471 | 565 | 467.6 | 500.0544 | 503 | 486.1 | 461 | 438.0145 | 482.7 | 508.762 | 488.9 | 432 | 446 | 495.9 | 500 | 539.9
| Mean |  | 399.8333 | 501.8333 | 501.8967 | 452.1667 | 457.3333 | 482.1667 | 455.8333 | 478.5 | 477.5 | 569.6667 | 479.9667 | 503.1653 | 504.5 | 483.615 | 470.5 | 451.5844 | 479.2167 | 489.2457 | 488.2667 | 441.8333 | 441.6667 | 493.6 | 495 | 543.4833
| Median |  | 397 | 504 | 499.015 | 449 | 456.5 | 483 | 454 | 480 | 478.5 | 569 | 477.55 | 503.4035 | 504.5 | 485.865 | 467.5 | 453.8208 | 480.75 | 489.436 | 487.85 | 443 | 444.5 | 494.3 | 495 | 545.1
| Std Dev. |  | 10.8151 | 9.7245 | 12.199 | 8.7044 | 3.3862 | 10.5909 | 4.2622 | 8.0685 | 4.6368 | 4.5461 | 9.8022 | 3.3255 | 8.3367 | 9.9118 | 9.3327 | 6.7972 | 4.9374 | 14.6698 | 5.9889 | 6.1779 | 10.5956 | 3.811 | 5.4772 | 8.0933
| Rel.Std.Dev. |  | 2.705% | 1.938% | 2.431% | 1.925% | 0.7404% | 2.197% | 0.935% | 1.686% | 0.9711% | 0.798% | 2.042% | 0.6609% | 1.652% | 2.05% | 1.984% | 1.505% | 1.03% | 2.998% | 1.227% | 1.398% | 2.399% | 0.7721% | 1.107% | 1.489%
| PDM3 |  | -16.81% | 4.407% | 4.42% | -5.927% | -4.852% | 0.315% | -5.164% | -0.4479% | -0.6559% | 18.52% | -0.1427% | 4.684% | 4.961% | 0.6163% | -2.112% | -6.048% | -0.2988% | 1.788% | 1.584% | -8.076% | -8.111% | 2.694% | 2.985% | 13.07%
| Z-Score (Absolute) |  | 2.63 | 0.67 | 0.68 | 0.94 | 0.77 | 0.04 | 0.82 | 0.08 | 0.12 | 2.87 | 0.04 | 0.72 | 0.76 | 0.08 | 0.34 | 0.96 | 0.06 | 0.27 | 0.23 | 1.27 | 1.28 | 0.41 | 0.45 | 2.02
| Analytical results for Pd in OREAS 611 (Indicative Value 196 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.03; 2.02: 2.11; 2.03: 2.12 | ppb
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 0
|  |  | 2.01: 15g; 2.02: 0.5g; 2.03: 01g | 0
| 1 | 1 | 2.01: < 10; 2.02: 200; 2.03: < 10 | 1
| 1 | 2 | 2.01: < 10; 2.02: 195; 2.03: < 10 | 1
| 1 | 3 | 2.01: < 10; 2.02: 190; 2.03: < 10 | 16
| 1 | 4 | 2.01: < 10; 2.02: 200; 2.03: < 10 | 195.8333
| 1 | 5 | 2.01: < 10; 2.02: 197; 2.03: < 10 | 17
| 1 | 6 | 2.01: < 10; 2.02: 193; 2.03: < 10
| Mean |  | 2.02: 195.8333
| Median |  | 2.02: 196
| Std Dev. |  | 2.02: 3.9707
| Rel.Std.Dev. |  | 2.02: 2.028%
| PDM3 |  | 2.02: 1.776e-13%
Z-Score (Absolute) | bar: 2.01 to 2.03
| Analytical results for Pr in OREAS 611 (Indicative Value 2.05 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | 2.01: Lab | 1
| No | No | 2.01: 2.04 | ppm
|  |  | 2.01: AR*MS | 2
|  |  | 2.01: 0.5g | 2
| 1 | 1 | 2.01: 1.9 | 1
| 1 | 2 | 2.01: 2.1 | 12
| 1 | 3 | 2.01: 2.1 | 16
| 1 | 4 | 2.01: 2.1 | 2.05
| 1 | 5 | 2.01: 2 | 18
| 1 | 6 | 2.01: 2.1
| Mean |  | 2.01: 2.05
| Median |  | 2.01: 2.1
| Std Dev. |  | 2.01: 0.083666
| Rel.Std.Dev. |  | 2.01: 4.081%
PDM3 | bar: 2.01
Z-Score (Absolute) | bar: 2.01
| Analytical results for Pt in OREAS 611 (Indicative Value < 1 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 2.01: Lab; 2.02: Lab; 2.03: Lab | 1
| No | No | 2.01: 2.03; 2.02: 2.11; 2.03: 2.12 | ppb
|  |  | 2.01: AR*MS; 2.02: AR*MS; 2.03: AR*OES/MS | 2
|  |  | 2.01: 15g; 2.02: 0.5g; 2.03: 01g | 2
| 1 | 1 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 1
| 1 | 2 | 2.01: 2; 2.02: < 1; 2.03: < 5 | 1
| 1 | 3 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 16
| 1 | 4 | 2.01: 2; 2.02: < 1; 2.03: < 5 | < 1
| 1 | 5 | 2.01: < 2; 2.02: < 1; 2.03: < 5 | 19
| 1 | 6 | 2.01: < 2; 2.02: < 1; 2.03: < 5
| Mean |  | 2.01: 2
| Median |  | 2.01: 2
| Std Dev. |  | 2.01: 0
Rel.Std.Dev. | bar: 2.01
PDM3
| Z-Score (Absolute) |  | 2.01: 0; 2.02: 0.67; 2.03: 0.94
| Analytical results for Rb in OREAS 611 (Certified Value 7.25 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.18 | 2.20 | 2.21 | 2.23 | [8 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | [8 empty] | 3
| 1 | 1 | 10.92 | 8.3 | 10.4 | 6.7 | 7.3 | 7.6 | 7.3 | 6.6 | 11.4 | 7.66 | 11.4933 | 6 | 6.8 | 7.7 | 7.1 | 6.8 | [8 empty] | 1
| 1 | 2 | 11.84 | 7.9 | 10.5 | 7.2 | 7.6 | 7.6 | 7.6 | 6.7 | 11.6 | 7.77 | 11.873 | 5.7 | 6.8 | 7.6 | 7.2 | 6.9 | [8 empty] | 41
| 1 | 3 | 12.58 | 8.1 | 10.6 | 6.9 | 7.5 | 7.6 | 7.4 | 6.7 | 11.9 | 7.91 | 11.6903 | 5.7 | 6.6 | 7.5 | 7.8 | 6.9 | [8 empty] | 16
| 1 | 4 | 11.83 | 8.5 | 10.4 | 7.1 | 7.8 | 7.6 | 7.6 | 6.6 | 11.1 | 7.99 | 11.6228 | 6 | 6.8 | 7.3 | 7.3 | 7.1 | [8 empty] | 7.2499
| 1 | 5 | 11.7 | 8.8 | 10.1 | 7.2 | 7.6 | 8.1 | 7.3 | 6.3 | 11.6 | 7.68 | 11.8375 | 6 | 6.8 | 7.5 | 7.7 | 7.3 | [8 empty] | 97
| 1 | 6 | 11.62 | 8.3 | 10.1 | 7.6 | 7.7 | 7.8 | 7.6 | 6.6 | 11.2 | 7.5 | 11.7243 | 5.6 | 6.9 | 7.3 | 7.6 | 6.6
| Mean |  | 11.7483 | 8.3167 | 10.35 | 7.1167 | 7.5833 | 7.7167 | 7.4667 | 6.5833 | 11.4667 | 7.7517 | 11.7069 | 5.8333 | 6.7833 | 7.4833 | 7.45 | 6.9333
| Median |  | 11.765 | 8.3 | 10.4 | 7.15 | 7.6 | 7.6 | 7.5 | 6.6 | 11.5 | 7.725 | 11.7073 | 5.85 | 6.8 | 7.5 | 7.45 | 6.9
| Std Dev. |  | 0.53135 | 0.31252 | 0.20736 | 0.30605 | 0.17224 | 0.20412 | 0.15055 | 0.1472 | 0.29439 | 0.17837 | 0.13994 | 0.18619 | 0.098319 | 0.16021 | 0.2881 | 0.24221
| Rel.Std.Dev. |  | 4.523% | 3.758% | 2.004% | 4.3% | 2.271% | 2.645% | 2.016% | 2.236% | 2.567% | 2.301% | 1.195% | 3.192% | 1.449% | 2.141% | 3.867% | 3.493%
| PDM3 |  | 62.05% | 14.71% | 42.76% | -1.837% | 4.6% | 6.439% | 2.99% | -9.194% | 58.16% | 6.922% | 61.48% | -19.54% | -6.435% | 3.22% | 2.761% | -4.366%
| Z-Score (Absolute) |  | 4.21 | 0.78 | 2.81 | 0.42 | 0.05 | 0.18 | 0.07 | 0.95 | 3.93 | 0.22 | 4.17 | 1.7 | 0.75 | 0.05 | 0.08 | 0.6
| Analytical results for Re in OREAS 611 (Indicative Value 1.83 ppb) | [25 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.18 | 2.20 | 2.23 | [10 empty] | ppb
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.2g | 0.25g | 0.25g | [10 empty] | 2
| 1 | 1 | < 50 | 2 | 2 | 4 | 1 | 2 | 2 | 2 | < 50 | 1 | < 50 | < 2 | 2 | < 5 | [10 empty] | 1
| 1 | 2 | < 50 | 2 | 2 | 1 | 2 | 2 | 1 | 1 | < 50 | 1 | < 50 | < 2 | 3 | < 5 | [10 empty] | 14
| 1 | 3 | < 50 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | < 50 | 2 | < 50 | < 2 | 3 | < 5 | [10 empty] | 16
| 1 | 4 | < 50 | 2 | 2 | 3 | 2 | 2 | 2 | 2 | < 50 | 2 | < 50 | < 2 | 3 | < 5 | [10 empty] | 1.8333
| 1 | 5 | < 50 | 2 | 2 | 2 | 2 | 2 | 2 | 1 | < 50 | 2 | < 50 | < 2 | 3 | < 5 | [10 empty] | 20
| 1 | 6 | < 50 | 1 | 2 | 2 | 2 | 2 | 2 | 1 | < 50 | 2 | < 50 | < 2 | 2 | < 5
| Mean |  |  | 1.8333 | 2 | 2.3333 | 1.8333 | 2 | 1.8333 | 1.5 |  | 1.6667 |  |  | 2.6667
| Median |  |  | 2 | 2 | 2 | 2 | 2 | 2 | 1.5 |  | 2 |  |  | 3
| Std Dev. |  |  | 0.40825 | 0 | 1.0328 | 0.40825 | 0 | 0.40825 | 0.54772 |  | 0.5164 |  |  | 0.5164
| Rel.Std.Dev. |  |  | 22.27% | 0% | 44.26% | 22.27% | 0% | 22.27% | 36.51% |  | 30.98% |  |  | 19.36%
| PDM3 |  |  | 1.776e-13% | 9.091% | 27.27% | 1.776e-13% | 9.091% | 1.776e-13% | -18.18% |  | -9.091% |  |  | 45.45%
| Z-Score (Absolute) |  | 37.22 | 0.27 | 0 | 0.54 | 0.27 | 0 | 0.27 | 0.81 | 37.22 | 0.54 | 37.22 | 1.62 | 1.08 | 0.81
| Analytical results for S in OREAS 611 (Certified Value 2.95 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 3.09 | 2.94 | 2.873 | 2.88 | 3.03 | 2.86 | 2.85 | 2.89 | 3.25 | 2.65 | 3.2824 | 2.24 | 2.75 | 2.98 | 2.6606 | 2.95 | 2.84 | 2.9 | 2.77 | 3.04 | 3.59 | 4.02 | 3.078 |  | 1
| 1 | 2 | 3.2 | 2.89 | 2.869 | 2.93 | 3.03 | 2.86 | 2.84 | 2.96 | 3.19 | 2.61 | 3.2992 | 2.22 | 2.74 | 2.92 | 2.6526 | 2.88 | 2.801 | 2.89 | 2.8 | 2.93 | 3.59 | 4.03 | 3.1 |  | 16
| 1 | 3 | 3.19 | 2.92 | 2.947 | 2.9 | 3.05 | 2.88 | 2.85 | 2.96 | 3.23 | 2.62 | 3.2593 | 2.17 | 2.84 | 2.99 | 2.6675 | 2.92 | 2.821 | 2.85 | 2.85 | 3.11 | 3.54 | 4.02 | 3.122 |  | 16
| 1 | 4 | 3.24 | 2.98 | 2.86 | 2.9 | 3.06 | 2.88 | 2.89 | 2.91 | 3.27 | 2.74 | 3.2252 | 2.19 | 2.87 | 3.12 | 2.6558 | 2.89 | 2.923 | 2.9 | 2.85 | 3.13 | 3.55 | 4.03 | 3.114 |  | 2.948
| 1 | 5 | 3.18 | 3 | 2.914 | 2.91 | 3.11 | 2.93 | 2.88 | 2.93 | 3.26 | 2.63 | 3.2285 | 2.17 | 2.82 | 3.11 | 2.6949 | 2.86 | 2.878 | 2.87 | 2.85 | 3.1 | 3.57 | 4.01 | 3.106 |  | 98
| 1 | 6 | 3.22 | 2.95 | 2.743 | 2.96 | 3.01 | 2.91 | 2.78 | 2.87 | 3.22 | 2.56 | 3.2896 | 2.12 | 2.81 | 2.97 | 2.5957 | 2.98 | 2.937 | 2.92 | 2.89 | 3.14 | 3.5 | 4.08 | 3.108
| Mean |  | 3.1867 | 2.9467 | 2.8677 | 2.9133 | 3.0483 | 2.8867 | 2.8483 | 2.92 | 3.2367 | 2.635 | 3.264 | 2.185 | 2.805 | 3.015 | 2.6545 | 2.9133 | 2.8667 | 2.8883 | 2.835 | 3.075 | 3.5567 | 4.0317 | 3.1047
| Median |  | 3.195 | 2.945 | 2.871 | 2.905 | 3.04 | 2.88 | 2.85 | 2.92 | 3.24 | 2.625 | 3.2708 | 2.18 | 2.815 | 2.985 | 2.6582 | 2.905 | 2.859 | 2.895 | 2.85 | 3.105 | 3.56 | 4.025 | 3.107
| Std Dev. |  | 0.052026 | 0.039833 | 0.069385 | 0.028048 | 0.034881 | 0.028048 | 0.038687 | 0.036878 | 0.029439 | 0.059582 | 0.031713 | 0.042308 | 0.050892 | 0.081179 | 0.032537 | 0.045461 | 0.055421 | 0.024833 | 0.042778 | 0.079183 | 0.034448 | 0.024833 | 0.015055
| Rel.Std.Dev. |  | 1.633% | 1.352% | 2.42% | 0.9627% | 1.144% | 0.9716% | 1.358% | 1.263% | 0.9096% | 2.261% | 0.9716% | 1.936% | 1.814% | 2.692% | 1.226% | 1.56% | 1.933% | 0.8598% | 1.509% | 2.575% | 0.9685% | 0.6159% | 0.4849%
| PDM3 |  | 8.097% | -0.0439% | -2.724% | -1.175% | 3.405% | -2.079% | -3.38% | -0.9485% | 9.793% | -10.62% | 10.72% | -25.88% | -4.849% | 2.274% | -9.954% | -1.175% | -2.758% | -2.023% | -3.832% | 4.309% | 20.65% | 36.76% | 5.316%
| Z-Score (Absolute) |  | 1.7 | 0.21 | 0.28 | 0 | 0.84 | 0.17 | 0.4 | 0.04 | 2.01 | 1.73 | 2.18 | 4.53 | 0.67 | 0.63 | 1.61 | 0 | 0.29 | 0.16 | 0.49 | 1.01 | 4 | 6.96 | 1.19
| Analytical results for Sb in OREAS 611 (Certified Value 324 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 0
| 1 | 1 | 256.51 | 338.12 | 320 | 311 | 351 | 320 | 335 | 326 | 300 | 339.21 | 326.6407 | 252 | 343.09 | 325 | 263.1596 | 235.26 | 308.929 | 283.85 | 319 | 264 | 350.2 | 320 | 345.5 |  | 1
| 1 | 2 | 268.48 | 344.59 | 340 | 315 | 346 | 331 | 333 | 329 | 303 | 331.82 | 325.0476 | 262 | 343.74 | 313 | 261.8188 | 231.3 | 314.197 | 278.05 | 358 | 259 | 355.7 | 320 | 345.3 |  | 43
| 1 | 3 | 264.55 | 345.28 | 331 | 313 | 351 | 326 | 332 | 330 | 303 | 329.71 | 328.9725 | 254 | 353.94 | 306 | 260.7359 | 232.77 | 317.21 | 275.84 | 334 | 276 | 348.5 | 320 | 332.1 |  | 16
| 1 | 4 | 267.22 | 352.61 | 325 | 314 | 357 | 324 | 341 | 327 | 303 | 337.23 | 325.5972 | 247 | 354.78 | 313 | 257.3772 | 233.84 | 317.492 | 266.53 | 332 | 273 | 355.5 | 320 | 341.9 |  | 324.4481
| 1 | 5 | 263.22 | 362.4 | 323 | 312 | 358 | 322 | 343 | 334 | 296 | 329.69 | 328.5525 | 247 | 348.53 | 319 | 274.4328 | 226.88 | 316.933 | 267.24 | 332 | 272 | 355.3 | 320 | 339.7 |  | 99
| 1 | 6 | 264.23 | 336.83 | 329 | 310 | 348 | 327 | 331 | 326 | 298 | 334.97 | 325.314 | 238 | 351.76 | 318 | 252.2535 | 238.69 | 328.727 | 278.46 | 339 | 277 | 367.8 | 320 | 341.5
| Mean |  | 264.035 | 346.6383 | 328 | 312.5 | 351.8333 | 325 | 335.8333 | 328.6667 | 300.5 | 333.7717 | 326.6874 | 250 | 349.3067 | 315.6667 | 261.6296 | 233.1233 | 317.248 | 274.995 | 335.6667 | 270.1667 | 355.5 | 320 | 341
| Median |  | 264.39 | 344.935 | 327 | 312.5 | 351 | 325 | 334 | 328 | 301.5 | 333.395 | 326.119 | 249.5 | 350.145 | 315.5 | 261.2773 | 233.305 | 317.0715 | 276.945 | 333 | 272.5 | 355.4 | 320 | 341.7
| Std Dev. |  | 4.1832 | 9.5802 | 7.0993 | 1.8708 | 4.7924 | 3.8987 | 4.9967 | 3.0768 | 3.0166 | 3.9984 | 1.701 | 8.0747 | 5.0534 | 6.5013 | 7.3904 | 3.9622 | 6.4876 | 6.8156 | 12.7854 | 7.1391 | 6.7596 | 0 | 4.9132
| Rel.Std.Dev. |  | 1.584% | 2.764% | 2.164% | 0.5987% | 1.362% | 1.2% | 1.488% | 0.9361% | 1.004% | 1.198% | 0.5207% | 3.23% | 1.447% | 2.06% | 2.825% | 1.7% | 2.045% | 2.478% | 3.809% | 2.642% | 1.901% | 0% | 1.441%
| PDM3 |  | -18.62% | 6.839% | 1.095% | -3.683% | 8.441% | 0.1701% | 3.509% | 1.3% | -7.381% | 2.874% | 0.6902% | -22.95% | 7.662% | -2.707% | -19.36% | -28.15% | -2.219% | -15.24% | 3.458% | -16.73% | 9.571% | -1.371% | 5.102%
| Z-Score (Absolute) |  | 2.57 | 0.91 | 0.13 | 0.53 | 1.13 | 0 | 0.46 | 0.15 | 1.03 | 0.37 | 0.07 | 3.16 | 1.02 | 0.39 | 2.67 | 3.87 | 0.33 | 2.11 | 0.45 | 2.31 | 1.29 | 0.21 | 0.67
| Analytical results for Sc in OREAS 611 (Certified Value 0.84 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 01g |  |  | 3
| 1 | 1 | < 1 | 1.6 | 1.1 | 0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 1 | 0.82 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.8 | 1.1 | 0.8 | 0.7 | 1 | 1.034 |  |  | 1
| 1 | 2 | 1 | 1.4 | 1 | 0.8 | 0.9 | 0.9 | 0.9 | 0.8 | 1 | 0.85 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.8 | 0.84 | 1.2 | 0.7 | < 0.5 | 1 | 1.08 |  |  | 44
| 1 | 3 | 1 | 1.7 | 0.8 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8 | 1 | 0.84 | < 1 | < 5 | 0.9 | 0.8 | < 5 | 0.7 | 0.84 | 1.1 | 0.8 | < 0.5 | 1 | 1.126 |  |  | 16
| 1 | 4 | 1 | 1.8 | 0.9 | 0.9 | 0.9 | 0.9 | 0.8 | 0.8 | 1 | 0.89 | < 1 | < 5 | 0.9 | 0.8 | < 5 | 0.8 | 0.84 | 1.1 | 0.8 | < 0.5 | 1 | 1.172 |  |  | 0.84326
| 1 | 5 | < 1 | 1.5 | 1 | 1 | 0.9 | 0.9 | 0.8 | 0.8 | 1 | 0.86 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.84 | 1.2 | 0.8 | < 0.5 | 1 | 1.118 |  |  | 100
| 1 | 6 | 1 | 1.5 | 1 | 0.9 | 0.8 | 0.9 | 0.9 | 0.8 | 1 | 0.77 | < 1 | < 5 | 0.8 | 0.8 | < 5 | 0.7 | 0.86 | 1.3 | 0.9 | < 0.5 | 0.9 | 1.164
| Mean |  | 1 | 1.5833 | 0.96667 | 0.9 | 0.88333 | 0.9 | 0.85 | 0.8 | 1 | 0.83833 |  |  | 0.83333 | 0.8 |  | 0.73333 | 0.83667 | 1.1667 | 0.8 | 0.7 | 0.98333 | 1.1157
| Median |  | 1 | 1.55 | 1 | 0.9 | 0.9 | 0.9 | 0.85 | 0.8 | 1 | 0.845 |  |  | 0.8 | 0.8 |  | 0.7 | 0.84 | 1.15 | 0.8 | 0.7 | 1 | 1.122
| Std Dev. |  | 0 | 0.1472 | 0.10328 | 0.063246 | 0.040825 | 0 | 0.054772 | 1.2162e-16 | 0 | 0.040702 |  |  | 0.05164 | 1.2162e-16 |  | 0.05164 | 0.019664 | 0.08165 | 0.063246 |  | 0.040825 | 0.052083
| Rel.Std.Dev. |  | 0% | 9.297% | 10.68% | 7.027% | 4.622% | 0% | 6.444% | 1.52e-14% | 0% | 4.855% |  |  | 6.197% | 1.52e-14% |  | 7.042% | 2.35% | 6.999% | 7.906% |  | 4.152% | 4.668%
| PDM3 |  | 18.59% | 87.76% | 14.63% | 6.728% | 4.752% | 6.728% | 0.7991% | -5.13% | 18.59% | -0.5845% |  |  | -1.177% | -5.13% |  | -13.04% | -0.7821% | 38.35% | -5.13% | -16.99% | 16.61% | 32.3%
| Z-Score (Absolute) |  | NA | 5.8 | 0.81 | 0.27 | 0.13 | 0.27 | 0.13 | 0.54 | NA | 0.23 | 2.97 | 13.22 | 0.27 | 0.54 | 13.22 | 1.08 | 0.24 | 2.43 | 0.54 | 4.38 | 0.94 | 2.01
Results from laboratories 2 & 10 were removed due to their 1 ppm readingresolution.
| Analytical results for Se in OREAS 611 (Certified Value 32.5 ppm) | [25 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.17 | 2.18 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | [5 empty] | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | [5 empty] | 1
| 1 | 1 | 29 | 35.5 | 39.5 | 33 | 37.1 | 36.3 | 34.8 | 34.8 | 31 | 33.9 | 31.3919 | 23.1 | 27.0935 | 25.2 | 37 | 28 | 40 | 20.8 | 32 | [5 empty] | 1
| 1 | 2 | 30 | 34.8 | 40.9 | 33.1 | 37.6 | 34.5 | 34.8 | 35.2 | 31 | 33.08 | 30.8619 | 25.7 | 29.4373 | 24 | 36 | 29 | 40 | 20.4 | 32 | [5 empty] | 45
| 1 | 3 | 30 | 35.5 | 39.7 | 33.7 | 37.5 | 35.9 | 35.8 | 37 | 30 | 33.23 | 31.9769 | 26.9 | 28.1891 | 23.1 | 37 | 29 | 40 | 20.9 | 32 | [5 empty] | 16
| 1 | 4 | 30 | 35.4 | 37.7 | 34.5 | 39 | 36.3 | 34.5 | 34.1 | 31 | 33.94 | 31.2805 | 30.5 | 26.0957 | 24.3 | 35 | 28 | 40 | 21.4 | 32 | [5 empty] | 32.5281
| 1 | 5 | 30 | 36.3 | 40 | 34.1 | 38.2 | 34.7 | 34.2 | 34.3 | 31 | 33.02 | 31.0965 | 28.2 | 27.7364 | 24.3 | 34 | 29 | 40 | 21.4 | 33 | [5 empty] | 101
| 1 | 6 | 31 | 35 | 38.5 | 35 | 30.8 | 35.9 | 32.8 | 36.7 | 30 | 32.68 | 32.2314 | 26.7 | 27.3432 | 24.2 | 35 | 30 | 40 | 17.9 | 33
| Mean |  | 30 | 35.4167 | 39.3833 | 33.9 | 36.7 | 35.6 | 34.4833 | 35.35 | 30.6667 | 33.3083 | 31.4732 | 26.85 | 27.6492 | 24.1833 | 35.6667 | 28.8333 | 40 | 20.4667 | 32.3333
| Median |  | 30 | 35.45 | 39.6 | 33.9 | 37.55 | 35.9 | 34.65 | 35 | 31 | 33.155 | 31.3362 | 26.8 | 27.5398 | 24.25 | 35.5 | 29 | 40 | 20.85 | 32
| Std Dev. |  | 0.63246 | 0.51929 | 1.1321 | 0.7874 | 2.9651 | 0.7975 | 0.98472 | 1.2276 | 0.5164 | 0.50701 | 0.52684 | 2.472 | 1.1236 | 0.67355 | 1.2111 | 0.75277 | 0 | 1.314 | 0.5164
| Rel.Std.Dev. |  | 2.108% | 1.466% | 2.875% | 2.323% | 8.079% | 2.24% | 2.856% | 3.473% | 1.684% | 1.522% | 1.674% | 9.207% | 4.064% | 2.785% | 3.395% | 2.611% | 0% | 6.42% | 1.597%
| PDM3 |  | -7.772% | 8.88% | 21.07% | 4.218% | 12.83% | 9.444% | 6.011% | 8.675% | -5.722% | 2.399% | -3.243% | -17.46% | -15% | -25.65% | 9.649% | -11.36% | 22.97% | -37.08% | -0.5987%
| Z-Score (Absolute) |  | 0.68 | 0.63 | 1.58 | 0.26 | 0.93 | 0.67 | 0.4 | 0.61 | 0.52 | 0.12 | 0.32 | 1.44 | 1.25 | 2.08 | 0.69 | 0.96 | NA | 2.98 | 0.12
Results from laboratory 22 were removed due to their 10 ppm reading resolution.
| Analytical results for Sm in OREAS 611 (Indicative Value 1.4 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | 2.02: Lab | [5 empty] | 1
| No | No | 2.02: 2.04 | [5 empty] | ppm
|  |  | 2.02: AR*MS | [5 empty] | 2
|  |  | 2.02: 0.5g | [5 empty] | 2
| 1 | 1 | 2.02: 1.5 | [5 empty] | 1
| 1 | 2 | 2.02: 1.5 | [5 empty] | 8
| 1 | 3 | 2.02: 1.3 | [5 empty] | 16
| 1 | 4 | 2.02: 1.6 | [5 empty] | 1.4
| 1 | 5 | 2.02: 1.4 | [5 empty] | 14
| 1 | 6 | 2.02: 1.1
| Mean |  | 2.02: 1.4
| Median |  | 2.02: 1.45
| Std Dev. |  | 2.02: 0.17889
| Rel.Std.Dev. |  | 2.02: 12.78%
| PDM3 |  | 2.02: 2.22e-14%
Z-Score (Absolute) | bar: 2.02
| Analytical results for Sn in OREAS 611 (Certified Value 29.5 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  |  | 2
| 1 | 1 | 28.9 | 32.9 | 28.6 | 28 | 30.6 | 30.4 | 30.6 | 29.8 | 26.7 | 31.14 | 28.7186 | 30.7 | 30.0186 | 26.5 | 30.32 | 31.8 | 28.6 | 27 | 30.9 | 28 | 26.26 |  |  |  | 1
| 1 | 2 | 30.4 | 33.1 | 29.5 | 28.2 | 31.5 | 30.3 | 30.9 | 30.4 | 26.7 | 30.3 | 28.8372 | 29.8 | 29.2053 | 26.6 | 30.948 | 31.1 | 28.6 | 27 | 31.1 | 28 | 26.67 |  |  |  | 18
| 1 | 3 | 29.8 | 32.3 | 29.4 | 28 | 31.4 | 30.8 | 31.8 | 30.3 | 26.6 | 30.23 | 29.1417 | 29.2 | 30.7364 | 26 | 29.809 | 30.6 | 29.4 | 29 | 31.1 | 28 | 27.08 |  |  |  | 16
| 1 | 4 | 30.2 | 34.4 | 28.6 | 28.3 | 31.8 | 30.5 | 30.3 | 30.1 | 26.7 | 31.19 | 29.1313 | 29.9 | 31.6816 | 26.7 | 30.338 | 29.3 | 28.1 | 29 | 31.3 | 29 | 27.49 |  |  |  | 29.5424
| 1 | 5 | 29.4 | 34.6 | 28.3 | 28.8 | 31.6 | 30.4 | 31.6 | 29.3 | 26.7 | 29.93 | 28.893 | 31 | 30.9955 | 26 | 30.514 | 29.6 | 28.4 | 28 | 31.7 | 28 | 27.9 |  |  |  | 102
| 1 | 6 | 30.3 | 33.2 | 28.5 | 29 | 30.3 | 30.5 | 30 | 30.1 | 26.8 | 30.32 | 28.6619 | 31 | 30.7412 | 26.2 | 30.233 | 30.9 | 29.1 | 28 | 31.2 | 29 | 26.31
| Mean |  | 29.8333 | 33.4167 | 28.8167 | 28.3833 | 31.2 | 30.4833 | 30.8667 | 30 | 26.7 | 30.5183 | 28.8973 | 30.2667 | 30.5631 | 26.3333 | 30.3603 | 30.55 | 28.7 | 28 | 31.2167 | 28.3333 | 26.9517
| Median |  | 30 | 33.15 | 28.6 | 28.25 | 31.45 | 30.45 | 30.75 | 30.1 | 26.7 | 30.31 | 28.8651 | 30.3 | 30.7388 | 26.35 | 30.329 | 30.75 | 28.6 | 28 | 31.15 | 28 | 26.875
| Std Dev. |  | 0.58878 | 0.89759 | 0.50365 | 0.42151 | 0.60332 | 0.17224 | 0.71461 | 0.4 | 0.063246 | 0.52036 | 0.2027 | 0.74207 | 0.85321 | 0.30768 | 0.37183 | 0.94393 | 0.47329 | 0.89443 | 0.27142 | 0.5164 | 0.65956
| Rel.Std.Dev. |  | 1.974% | 2.686% | 1.748% | 1.485% | 1.934% | 0.565% | 2.315% | 1.333% | 0.2369% | 1.705% | 0.7015% | 2.452% | 2.792% | 1.168% | 1.225% | 3.09% | 1.649% | 3.194% | 0.8695% | 1.823% | 2.447%
| PDM3 |  | 0.9847% | 13.11% | -2.457% | -3.923% | 5.611% | 3.185% | 4.483% | 1.549% | -9.621% | 3.303% | -2.184% | 2.452% | 3.455% | -10.86% | 2.769% | 3.411% | -2.852% | -5.221% | 5.667% | -4.093% | -8.77%
| Z-Score (Absolute) |  | 0.09 | 1.95 | 0.67 | 0.92 | 0.68 | 0.28 | 0.49 | 0 | 1.88 | 0.3 | 0.63 | 0.15 | 0.32 | 2.09 | 0.21 | 0.31 | 0.74 | 1.14 | 0.69 | 0.95 | 1.74
| Analytical results for Sr in OREAS 611 (Certified Value 38.4 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 1
| 1 | 1 | 46.5 | 36.1 | 43.9 | 35 | 37 | 37.7 | 38.4 | 39.4 | 51.6 | 38.37 | 56.2786 | 22 | 39.8 | 34 | 49.26 | 30.3 | 34.9 | 40.4 | 37.8 | 36.9 | 31 | 27 | 64.88 |  | 1
| 1 | 2 | 49.7 | 36.1 | 49.5 | 36.5 | 39.2 | 38.9 | 38.7 | 39.6 | 51.6 | 39.01 | 55.1797 | 22 | 38.2 | 33.5 | 48.5621 | 30.8 | 35.43 | 39.8 | 37.7 | 39.1 | 31 | 29 | 65.09 |  | 19
| 1 | 3 | 52 | 36.5 | 48.9 | 36.2 | 38.3 | 38.3 | 37.7 | 40.1 | 51.6 | 39.03 | 54.9921 | 22 | 41.3 | 34.8 | 47.7722 | 28.8 | 38.06 | 40.7 | 40.1 | 39.1 | 32 | 28 | 65.3 |  | 16
| 1 | 4 | 49.9 | 37.2 | 49 | 36.9 | 39 | 38.9 | 38.3 | 38.9 | 51.4 | 39.6 | 53.7701 | 21 | 41.4 | 35.1 | 47.7356 | 30.5 | 34.54 | 39 | 39.6 | 38.9 | 31 | 28 | 65.51 |  | 38.4178
| 1 | 5 | 48.9 | 39.8 | 45.7 | 37.7 | 38.4 | 41.5 | 36.7 | 37.9 | 51.7 | 38.13 | 55.6755 | 21 | 41 | 34 | 49.4118 | 30 | 36.56 | 38.1 | 40 | 39.2 | 32 | 28 | 65.72 |  | 103
| 1 | 6 | 49 | 37.1 | 49.5 | 35.9 | 39.3 | 40.1 | 35.9 | 38.4 | 51.3 | 37.5 | 54.3942 | 20 | 40.3 | 34.1 | 45.7269 | 30.4 | 34.65 | 39.5 | 40.5 | 38.5 | 32 | 28 | 65.93
| Mean |  | 49.3333 | 37.1333 | 47.75 | 36.3667 | 38.5333 | 39.2333 | 37.6167 | 39.05 | 51.5333 | 38.6067 | 55.0484 | 21.3333 | 40.3333 | 34.25 | 48.0781 | 30.1333 | 35.69 | 39.5833 | 39.2833 | 38.6167 | 31.5 | 28 | 65.405
| Median |  | 49.35 | 36.8 | 48.95 | 36.35 | 38.7 | 38.9 | 38 | 39.15 | 51.6 | 38.69 | 55.0859 | 21.5 | 40.65 | 34.05 | 48.1672 | 30.35 | 35.165 | 39.65 | 39.8 | 39 | 31.5 | 28 | 65.405
| Std Dev. |  | 1.7829 | 1.3895 | 2.3679 | 0.91579 | 0.85713 | 1.3663 | 1.0998 | 0.81179 | 0.15055 | 0.7532 | 0.89319 | 0.8165 | 1.2127 | 0.58907 | 1.3527 | 0.70333 | 1.3762 | 0.94956 | 1.2222 | 0.87731 | 0.54772 | 0.63246 | 0.39287
| Rel.Std.Dev. |  | 3.614% | 3.742% | 4.959% | 2.518% | 2.224% | 3.482% | 2.924% | 2.079% | 0.2921% | 1.951% | 1.623% | 3.827% | 3.007% | 1.72% | 2.814% | 2.334% | 3.856% | 2.399% | 3.111% | 2.272% | 1.739% | 2.259% | 0.6007%
| PDM3 |  | 28.41% | -3.343% | 24.29% | -5.339% | 0.3007% | 2.123% | -2.085% | 1.646% | 34.14% | 0.4916% | 43.29% | -44.47% | 4.986% | -10.85% | 25.15% | -21.56% | -7.1% | 3.034% | 2.253% | 0.5177% | -18.01% | -27.12% | 70.25%
| Z-Score (Absolute) |  | 2.47 | 0.34 | 2.1 | 0.52 | 0.02 | 0.14 | 0.23 | 0.1 | 2.98 | 0 | 3.79 | 3.98 | 0.4 | 1.01 | 2.18 | 1.95 | 0.67 | 0.22 | 0.15 | 0 | 1.64 | 2.45 | 6.17
| Analytical results for Ta in OREAS 611 (Indicative Value < 0.01 ppm) | [25 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [7 empty] | 3
| 1 | 1 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19391 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19427 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 9
| 1 | 3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19486 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 16
| 1 | 4 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19523 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19364 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1 | [7 empty] | 15
| 1 | 6 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.19353 | < 0.01 | < 10 | < 0.3 | < 0.05 | < 0.05 | < 0.1
| Mean |  | 2.12: 0.19424
| Median |  | 2.12: 0.19409
| Std Dev. |  | 2.12: 0.00068341
| Rel.Std.Dev. |  | 2.12: 0.3518%
PDM3
| Z-Score (Absolute) |  | 0 | 0 | 0 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0 | 0.67 | 5.71 | 0.67 | 167.73 | 4.21 | 0 | 0 | 0.84
| Analytical results for Tb in OREAS 611 (Indicative Value 0.17 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab | [7 empty] | 1
| No | No | 2.02: 2.04; 2.03: 2.18; 2.04: 2.20; 2.05: 2.21 | [7 empty] | ppm
|  |  | 2.02: AR*MS; 2.03: AR*OES/MS; 2.04: AR*OES/MS; 2.05: AR*OES/MS | [7 empty] | 2
|  |  | 2.02: 0.5g; 2.03: 0.2g; 2.04: 0.25g; 2.05: 0.25g | [7 empty] | 2
| 1 | 1 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.22; 2.05: 0.18 | [7 empty] | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.19; 2.05: 0.17 | [7 empty] | 10
| 1 | 3 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.21; 2.05: 0.17 | [7 empty] | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.19; 2.05: 0.17 | [7 empty] | 0.16917
| 1 | 5 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.21; 2.05: 0.17 | [7 empty] | 16
| 1 | 6 | 2.02: 0.2; 2.03: 0.1; 2.04: 0.21; 2.05: 0.17
| Mean |  | 2.02: 0.2; 2.03: 0.1; 2.04: 0.205; 2.05: 0.17167
| Median |  | 2.02: 0.2; 2.03: 0.1; 2.04: 0.21; 2.05: 0.17
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 1.5202e-17; 2.04: 0.012247; 2.05: 0.0040825
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 1.52e-14%; 2.04: 5.974%; 2.05: 2.378%
| PDM3 |  | 2.02: 18.23%; 2.03: -40.89%; 2.04: 21.18%; 2.05: 1.478%
| Z-Score (Absolute) |  | 2.02: 0.57; 2.03: 3.47; 2.04: 0.78; 2.05: 0.57
| Analytical results for Te in OREAS 611 (Certified Value 49 ppm) | [25 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.24 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | [5 empty] | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | [5 empty] | 2
| 1 | 1 | 45.75 | 53.19 | 43.6 | 46.9 | 52.9 | 47.7 | 52.2 | 50 | 46.1 | 48.46 | 51.4266 | 47.38 | 52 | 52.0682 | 50.33 | 51.93 | 45.51 | 45.7 | 35 | [5 empty] | 1
| 1 | 2 | 46.84 | 54.41 | 46.4 | 47.1 | 52.8 | 46.8 | 51.1 | 50.6 | 47.5 | 46.82 | 50.3021 | 46.14 | 50.6 | 50.8 | 48.43 | 51.51 | 45.64 | 45.6 | 35 | [5 empty] | 34
| 1 | 3 | 46.94 | 53.96 | 46.2 | 47.3 | 53.8 | 47.8 | 52.6 | 49.8 | 46.3 | 46.33 | 49.8703 | 48.58 | 51.3 | 51.3355 | 47.47 | 50.13 | 46.03 | 44.6 | 35 | [5 empty] | 16
| 1 | 4 | 46.88 | 54.83 | 43.7 | 48 | 54.5 | 46.8 | 50.4 | 49.3 | 47.3 | 47.04 | 49.8411 | 47.16 | 51.2 | 50.8034 | 49.53 | 49.15 | 45.66 | 46 | 35 | [5 empty] | 48.9911
| 1 | 5 | 46.73 | 57.8 | 44.1 | 48.9 | 54.3 | 49.1 | 51.9 | 48.7 | 47 | 45.55 | 51.559 | 49.04 | 51.1 | 52.4599 | 47.87 | 48.92 | 46.48 | 46.5 | 35 | [5 empty] | 104
| 1 | 6 | 47.67 | 53.44 | 44.8 | 49.6 | 54.2 | 48.9 | 50 | 50 | 46.8 | 46.28 | 51.7925 | 48.19 | 50.8 | 51.0177 | 48.73 | 51.56 | 46.43 | 55.1 | 35
| Mean |  | 46.8017 | 54.605 | 44.8 | 47.9667 | 53.75 | 47.85 | 51.3667 | 49.7333 | 46.8333 | 46.7467 | 50.7986 | 47.7483 | 51.1667 | 51.4141 | 48.7267 | 50.5333 | 45.9583 | 47.25 | 35
| Median |  | 46.86 | 54.185 | 44.45 | 47.65 | 54 | 47.75 | 51.5 | 49.9 | 46.9 | 46.575 | 50.8643 | 47.785 | 51.15 | 51.1766 | 48.58 | 50.82 | 45.845 | 45.85 | 35
| Std Dev. |  | 0.61512 | 1.6776 | 1.2377 | 1.0838 | 0.73417 | 0.98944 | 1.0367 | 0.65625 | 0.55015 | 0.98492 | 0.89281 | 1.0605 | 0.48442 | 0.69785 | 1.0599 | 1.3144 | 0.42216 | 3.896 | 0
| Rel.Std.Dev. |  | 1.314% | 3.072% | 2.763% | 2.26% | 1.366% | 2.068% | 2.018% | 1.32% | 1.175% | 2.107% | 1.758% | 2.221% | 0.9468% | 1.357% | 2.175% | 2.601% | 0.9186% | 8.246% | 0%
| PDM3 |  | -4.469% | 11.46% | -8.555% | -2.091% | 9.714% | -2.329% | 4.849% | 1.515% | -4.404% | -4.581% | 3.689% | -2.537% | 4.441% | 4.946% | -0.5398% | 3.148% | -6.191% | -3.554% | -28.56%
| Z-Score (Absolute) |  | 0.39 | 2.23 | 1.06 | 0 | 1.94 | 0.04 | 1.14 | 0.59 | 0.38 | 0.41 | 0.95 | 0.07 | 1.07 | 1.16 | 0.26 | 0.86 | 0.67 | 0.24 | 4.35
| Analytical results for Th in OREAS 611 (Certified Value 3.14 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.25 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 01g |  |  |  |  | 3
| 1 | 1 | 4.85 | 3.3 | 3.3 | 3 | 2.7 | 3 | 3 | 2.8 | 5.37 | 3.4 | 4.598 | 6 | 3 | 3.2 | < 10 | 2.79 | 3.5 | 3.2 | 2.8 | < 2 |  |  |  |  | 1
| 1 | 2 | 5.02 | 3.2 | 3.6 | 3.1 | 2.9 | 3 | 3 | 2.8 | 5.37 | 3.432 | 4.3956 | 6 | 3.1 | 3.2 | < 10 | 2.77 | 3.4 | 3.2 | 2.9 | < 2 |  |  |  |  | 35
| 1 | 3 | 5.33 | 3.2 | 3.8 | 3 | 2.8 | 2.9 | 3 | 2.8 | 5.37 | 3.419 | 4.4354 | 5 | 3.5 | 3.2 | < 10 | 2.73 | 3.5 | 3.4 | 3 | < 2 |  |  |  |  | 16
| 1 | 4 | 5.13 | 3.4 | 3.7 | 2.9 | 2.9 | 3 | 3.1 | 2.9 | 5.37 | 3.427 | 4.4773 | 6 | 3.6 | 3.3 | < 10 | 2.79 | 3.4 | 3.4 | 3.1 | < 2 |  |  |  |  | 3.1383
| 1 | 5 | 4.96 | 3.5 | 3.6 | 3 | 2.9 | 3.2 | 3 | 2.8 | 5.44 | 3.435 | 4.6432 | 5 | 3.3 | 3.2 | < 10 | 2.73 | 3.3 | 3.4 | 3.1 | < 2 |  |  |  |  | 105
| 1 | 6 | 4.82 | 3.2 | 3.6 | 2.9 | 3.3 | 3 | 2.9 | 2.7 | 5.23 | 3.414 | 4.555 | 5 | 3.3 | 3.1 | < 10 | 2.84 | 3.4 | 3.4 | 2.8 | < 2
| Mean |  | 5.0183 | 3.3 | 3.6 | 2.9833 | 2.9167 | 3.0167 | 3 | 2.8 | 5.3583 | 3.4212 | 4.5174 | 5.5 | 3.3 | 3.2 |  | 2.775 | 3.4167 | 3.3333 | 2.95
| Median |  | 4.99 | 3.25 | 3.6 | 3 | 2.9 | 3 | 3 | 2.8 | 5.37 | 3.423 | 4.5161 | 5.5 | 3.3 | 3.2 |  | 2.78 | 3.4 | 3.4 | 2.95
| Std Dev. |  | 0.1901 | 0.12649 | 0.16733 | 0.075277 | 0.20412 | 0.098319 | 0.063246 | 0.063246 | 0.068823 | 0.013014 | 0.096853 | 0.54772 | 0.22804 | 0.063246 |  | 0.041833 | 0.075277 | 0.10328 | 0.13784
| Rel.Std.Dev. |  | 3.788% | 3.833% | 4.648% | 2.523% | 6.999% | 3.259% | 2.108% | 2.259% | 1.284% | 0.3804% | 2.144% | 9.959% | 6.91% | 1.976% |  | 1.507% | 2.203% | 3.098% | 4.673%
| PDM3 |  | 59.91% | 5.153% | 14.71% | -4.938% | -7.062% | -3.876% | -4.407% | -10.78% | 70.74% | 9.013% | 43.94% | 75.25% | 5.153% | 1.966% |  | -11.58% | 8.87% | 6.215% | -6%
| Z-Score (Absolute) |  | 3.66 | 0 | 0.64 | 0.67 | 0.82 | 0.6 | 0.64 | 1.06 | 4.38 | 0.26 | 2.59 | NA | 0 | 0.21 | 3.62 | 1.12 | 0.25 | 0.07 | 0.75 | 4.9
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 611 (Indicative Value 0.01 wt.%) | [25 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | wt.%
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 3
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 3
| 1 | 1 | 0.0162 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.0212 | 0.0101 | 0.021184 | 0.008 | 0.01 | 0.008 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.008 | 0.01 | 0.04461 |  | 1
| 1 | 2 | 0.0176 | 0.01 | 0.02 | 0.01 | 0.011 | 0.01 | 0.01 | 0.011 | 0.0208 | 0.01 | 0.021735 | 0.008 | 0.009 | 0.007 | < 0.01 | < 0.01 | < 0.01 | 0.01 | < 0.01 | 0.01 | 0.008 | 0.01 | 0.04824 |  | 11
| 1 | 3 | 0.0201 | 0.01 | 0.02 | 0.01 | 0.011 | 0.01 | 0.01 | 0.011 | 0.0217 | 0.0099 | 0.021086 | 0.008 | 0.01 | 0.008 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.008 | 0.01 | 0.04687 |  | 16
| 1 | 4 | 0.0182 | 0.01 | 0.02 | 0.01 | 0.011 | 0.01 | 0.011 | 0.01 | 0.0212 | 0.0101 | 0.021322 | 0.008 | 0.01 | 0.008 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.008 | 0.01 | 0.0455 |  | 0.010127
| 1 | 5 | 0.0172 | 0.011 | 0.02 | 0.01 | 0.011 | 0.011 | 0.01 | 0.01 | 0.0217 | 0.01 | 0.021629 | 0.008 | 0.01 | 0.008 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.009 | 0.01 | 0.04913 |  | 17
| 1 | 6 | 0.017 | 0.01 | 0.02 | 0.01 | 0.011 | 0.011 | 0.009 | 0.01 | 0.021 | 0.009 | 0.021829 | 0.007 | 0.01 | 0.007 | < 0.01 | < 0.01 | < 0.01 | 0.01 | 0.01 | 0.01 | 0.008 | 0.01 | 0.05276
| Mean |  | 0.017717 | 0.010167 | 0.02 | 0.01 | 0.010833 | 0.010333 | 0.01 | 0.010333 | 0.021267 | 0.00985 | 0.021464 | 0.0078333 | 0.0098333 | 0.0076667 |  |  |  | 0.01 | 0.01 | 0.01 | 0.0081667 | 0.01 | 0.047852
| Median |  | 0.0174 | 0.01 | 0.02 | 0.01 | 0.011 | 0.01 | 0.01 | 0.01 | 0.0212 | 0.01 | 0.021476 | 0.008 | 0.01 | 0.008 |  |  |  | 0.01 | 0.01 | 0.01 | 0.008 | 0.01 | 0.047555
| Std Dev. |  | 0.0013423 | 0.00040825 | 0 | 0 | 0.00040825 | 0.0005164 | 0.00063246 | 0.0005164 | 0.00036697 | 0.00042308 | 0.00030847 | 0.00040825 | 0.00040825 | 0.0005164 |  |  |  | 0 | 0 | 0 | 0.00040825 | 0 | 0.0029284
| Rel.Std.Dev. |  | 7.576% | 4.016% | 0% | 0% | 3.768% | 4.997% | 6.325% | 4.997% | 1.726% | 4.295% | 1.437% | 5.212% | 4.152% | 6.736% |  |  |  | 0% | 0% | 0% | 4.999% | 0% | 6.12%
| PDM3 |  | 74.95% | 0.395% | 97.5% | -1.251% | 6.978% | 2.041% | -1.251% | 2.041% | 110% | -2.732% | 112% | -22.65% | -2.897% | -24.29% |  |  |  | -1.251% | -1.251% | -1.251% | -19.35% | -1.251% | 372.5%
| Z-Score (Absolute) |  | 6.24 | 0.13 | 8.09 | 0 | 0.67 | 0.27 | 0 | 0.27 | 9.12 | 0.12 | 9.28 | 1.75 | 0.13 | 1.89 | 4.05 | 4.05 | 4.05 | 0 | 0.67 | 0 | 1.48 | 0 | 30.63
| Analytical results for Tl in OREAS 611 (Certified Value 1.81 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | 2.25 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 01g |  |  |  |  | 3
| 1 | 1 | 1.79 | 1.93 | 1.87 | 1.78 | 1.83 | 1.79 | 1.75 | 1.77 | 1.83 | 1.84 | 1.8481 | < 5 | 1.58 | 1.47 | < 5 | 1.47 | 1.81 | 1.71 | 2 | < 5 |  |  |  |  | 1
| 1 | 2 | 1.82 | 1.92 | 1.93 | 1.85 | 1.77 | 1.75 | 1.75 | 1.81 | 1.83 | 1.79 | 1.8638 | < 5 | 1.67 | 1.51 | < 5 | 1.45 | 1.82 | 1.72 | 2 | < 5 |  |  |  |  | 37
| 1 | 3 | 1.88 | 1.97 | 1.95 | 1.82 | 1.86 | 1.72 | 1.75 | 1.73 | 1.8 | 1.76 | 1.9138 | < 5 | 1.86 | 1.57 | < 5 | 1.44 | 1.86 | 1.76 | 2 | < 5 |  |  |  |  | 16
| 1 | 4 | 1.85 | 2 | 1.9 | 1.72 | 1.81 | 1.77 | 1.71 | 1.79 | 1.86 | 1.82 | 1.8062 | < 5 | 1.88 | 1.49 | < 5 | 1.44 | 1.76 | 1.67 | 2 | < 5 |  |  |  |  | 1.8124
| 1 | 5 | 1.86 | 2.04 | 1.87 | 1.78 | 1.83 | 1.84 | 1.76 | 1.86 | 1.87 | 1.77 | 1.9447 | < 5 | 1.74 | 1.49 | < 5 | 1.41 | 1.74 | 1.78 | 2 | < 5 |  |  |  |  | 106
| 1 | 6 | 1.87 | 1.95 | 1.87 | 1.68 | 1.8 | 1.77 | 1.7 | 1.81 | 1.83 | 1.74 | 1.8679 | < 5 | 1.71 | 1.48 | < 5 | 1.5 | 1.86 | 1.7 | 2 | < 5
| Mean |  | 1.845 | 1.9683 | 1.8983 | 1.7717 | 1.8167 | 1.7733 | 1.7367 | 1.795 | 1.8367 | 1.7867 | 1.8741 |  | 1.74 | 1.5017 |  | 1.4517 | 1.8083 | 1.7233 | 2
| Median |  | 1.855 | 1.96 | 1.885 | 1.78 | 1.82 | 1.77 | 1.75 | 1.8 | 1.83 | 1.78 | 1.8658 |  | 1.725 | 1.49 |  | 1.445 | 1.815 | 1.715 | 2
| Std Dev. |  | 0.033912 | 0.04535 | 0.034881 | 0.062743 | 0.030768 | 0.040332 | 0.025033 | 0.043704 | 0.025033 | 0.037771 | 0.048967 |  | 0.11437 | 0.036009 |  | 0.030605 | 0.049967 | 0.040332 | 0
| Rel.Std.Dev. |  | 1.838% | 2.304% | 1.837% | 3.541% | 1.694% | 2.274% | 1.441% | 2.435% | 1.363% | 2.114% | 2.613% |  | 6.573% | 2.398% |  | 2.108% | 2.763% | 2.34% | 0%
| PDM3 |  | 1.797% | 8.602% | 4.739% | -2.249% | 0.2335% | -2.157% | -4.18% | -0.962% | 1.337% | -1.422% | 3.402% |  | -3.997% | -17.15% |  | -19.91% | -0.2263% | -4.916% | 10.35%
| Z-Score (Absolute) |  | 0.36 | 1.58 | 0.89 | 0.36 | 0.08 | 0.35 | 0.71 | 0.13 | 0.28 | 0.21 | 0.65 | 6.83 | 0.67 | 3.03 | 6.83 | 3.52 | 0 | 0.84 | NA | 6.83
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Tm in OREAS 611 (Indicative Value < 100 ppb) | [25 empty] | Indicative
Tm | Round | Replicate | 2.02: Lab |  |  |  |  | 1
| No | No | 2.02: 2.04 |  |  |  |  | ppb
|  |  | 2.02: AR*MS |  |  |  |  | 0
|  |  | 2.02: 0.5g |  |  |  |  | 0
| 1 | 1 | 2.02: < 100 |  |  |  |  | 1
| 1 | 2 | 2.02: < 100 |  |  |  |  | 6
| 1 | 3 | 2.02: < 100 |  |  |  |  | 16
| 1 | 4 | 2.02: < 100 |  |  |  |  | < 100
| 1 | 5 | 2.02: < 100 |  |  |  |  | 18
| 1 | 6 | 2.02: < 100
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 2.02
| Analytical results for U in OREAS 611 (Certified Value 1.12 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.14 | 2.15 | 2.17 | 2.18 | 2.20 | 2.21 | 2.23 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | [6 empty] | 3
| 1 | 1 | 1.17 | 1.2 | 1.2 | 1.04 | 1.03 | 1.09 | 1.09 | 1.02 | 1.22 | 1.177 | 1.1977 | 0.92 | 0.94 | < 10 | 0.82 | 1.16 | 1.1 | 1.1 | [6 empty] | 1
| 1 | 2 | 1.21 | 1.3 | 1.3 | 1.02 | 1.06 | 1.06 | 1.11 | 1.04 | 1.24 | 1.148 | 1.1811 | 0.98 | 0.96 | < 10 | 0.82 | 1.11 | 1.08 | 1.1 | [6 empty] | 39
| 1 | 3 | 1.25 | 1.3 | 1.3 | 1 | 1.05 | 1.06 | 1.09 | 1.03 | 1.2 | 1.138 | 1.2025 | 1.1 | 0.96 | < 10 | 0.82 | 1.17 | 1.1 | 1.2 | [6 empty] | 16
| 1 | 4 | 1.24 | 1.3 | 1.2 | 1 | 1.06 | 1.06 | 1.09 | 1.05 | 1.2 | 1.159 | 1.1584 | 1.12 | 1 | < 10 | 0.83 | 1.09 | 1.11 | 1.2 | [6 empty] | 1.1214
| 1 | 5 | 1.2 | 1.3 | 1.2 | 1.05 | 1.07 | 1.11 | 1.08 | 1.05 | 1.2 | 1.142 | 1.216 | 1.11 | 0.95 | < 10 | 0.81 | 1.1 | 1.11 | 1.2 | [6 empty] | 107
| 1 | 6 | 1.22 | 1.3 | 1.2 | 1.11 | 1.11 | 1.09 | 1.02 | 1.03 | 1.21 | 1.141 | 1.1662 | 1.05 | 0.95 | < 10 | 0.85 | 1.12 | 1.08 | 1
| Mean |  | 1.215 | 1.2833 | 1.2333 | 1.0367 | 1.0633 | 1.0783 | 1.08 | 1.0367 | 1.2117 | 1.1508 | 1.187 | 1.0467 | 0.96 |  | 0.825 | 1.125 | 1.0967 | 1.1333
| Median |  | 1.215 | 1.3 | 1.2 | 1.03 | 1.06 | 1.075 | 1.09 | 1.035 | 1.205 | 1.145 | 1.1894 | 1.075 | 0.955 |  | 0.82 | 1.115 | 1.1 | 1.15
| Std Dev. |  | 0.02881 | 0.040825 | 0.05164 | 0.041312 | 0.026583 | 0.02137 | 0.030984 | 0.012111 | 0.016021 | 0.014825 | 0.022278 | 0.080911 | 0.020976 |  | 0.013784 | 0.032711 | 0.013663 | 0.08165
| Rel.Std.Dev. |  | 2.371% | 3.181% | 4.187% | 3.985% | 2.5% | 1.982% | 2.869% | 1.168% | 1.322% | 1.288% | 1.877% | 7.73% | 2.185% |  | 1.671% | 2.908% | 1.246% | 7.204%
| PDM3 |  | 8.35% | 14.44% | 9.985% | -7.553% | -5.175% | -3.837% | -3.689% | -7.553% | 8.053% | 2.628% | 5.851% | -6.661% | -14.39% |  | -26.43% | 0.3243% | -2.202% | 1.067%
| Z-Score (Absolute) |  | 0.95 | 1.57 | 1.11 | 0.67 | 0.43 | 0.3 | 0.28 | 0.67 | 0.92 | 0.36 | 0.69 | 0.58 | 1.37 | 35.36 | 2.6 | 0.13 | 0.13 | 0.2
| Analytical results for V in OREAS 611 (Certified Value 12.7 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  | 2
| 1 | 1 | 14 | 13 | 14 | 11 | 12 | 12 | 12 | 12 | 15 | 12.1 | 10.5611 | 13 | 11 | 13.7 | 12.2496 | 10.3 | 11.963 | 14 | 13 | 14 | 12 | 12 | 16.49 |  | 1
| 1 | 2 | 15 | 12 | 14 | 12 | 13 | 12 | 12 | 12 | 15 | 12.8 | 10.6374 | 13 | 10 | 13.7 | 11.9275 | 10.3 | 12.04 | 14 | 12 | 14 | 12 | 12 | 17 |  | 40
| 1 | 3 | 15 | 13 | 14 | 12 | 13 | 12 | 12 | 12 | 15 | 12.9 | 10.7506 | 13 | 14 | 14 | 12.1644 | 10 | 12.281 | 14 | 13 | 15 | 13 | 12 | 16.51 |  | 16
| 1 | 4 | 15 | 13 | 14 | 12 | 13 | 12 | 12 | 12 | 15 | 13 | 10.8871 | 13 | 13 | 14 | 12.2037 | 10.3 | 12.117 | 14 | 12 | 14 | 12 | 12 | 17.02 |  | 12.6768
| 1 | 5 | 15 | 13 | 14 | 12 | 13 | 13 | 12 | 12 | 15 | 12.7 | 10.4448 | 13 | 11 | 13.9 | 12.2712 | 10 | 12.495 | 14 | 13 | 14 | 13 | 12 | 16.53 |  | 108
| 1 | 6 | 15 | 13 | 14 | 12 | 12 | 12 | 12 | 12 | 15 | 11.8 | 10.4507 | 13 | 11 | 13.9 | 11.6992 | 10 | 12.715 | 14 | 12 | 14 | 13 | 12 | 17.04
| Mean |  | 14.8333 | 12.8333 | 14 | 11.8333 | 12.6667 | 12.1667 | 12 | 12 | 15 | 12.55 | 10.622 | 13 | 11.6667 | 13.8667 | 12.0859 | 10.15 | 12.2685 | 14 | 12.5 | 14.1667 | 12.5 | 12 | 16.765
| Median |  | 15 | 13 | 14 | 12 | 13 | 12 | 12 | 12 | 15 | 12.75 | 10.5993 | 13 | 11 | 13.9 | 12.184 | 10.15 | 12.199 | 14 | 12.5 | 14 | 12.5 | 12 | 16.765
| Std Dev. |  | 0.40825 | 0.40825 | 0 | 0.40825 | 0.5164 | 0.40825 | 0 | 0 | 0 | 0.48477 | 0.17406 | 0 | 1.5055 | 0.13663 | 0.22618 | 0.16432 | 0.28959 | 0 | 0.54772 | 0.40825 | 0.54772 | 0 | 0.27991
| Rel.Std.Dev. |  | 2.752% | 3.181% | 0% | 3.45% | 4.077% | 3.355% | 0% | 0% | 0% | 3.863% | 1.639% | 0% | 12.9% | 0.9853% | 1.871% | 1.619% | 2.36% | 0% | 4.382% | 2.882% | 4.382% | 0% | 1.67%
| PDM3 |  | 17.01% | 1.235% | 10.44% | -6.653% | -0.07966% | -4.024% | -5.339% | -5.339% | 18.33% | -1% | -16.21% | 2.55% | -7.968% | 9.386% | -4.661% | -19.93% | -3.221% | 10.44% | -1.394% | 11.75% | -1.394% | -5.339% | 32.25%
| Z-Score (Absolute) |  | 3.15 | 0.45 | 2.02 | 0.9 | 0.22 | 0.45 | 0.67 | 0.67 | 3.37 | 0.07 | 2.53 | 0.67 | 1.12 | 1.84 | 0.56 | 3.17 | 0.31 | 2.02 | 0 | 2.25 | 0 | 0.67 | 5.75
| Analytical results for W in OREAS 611 (Certified Value 4.31 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.25 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 01g |  |  | 2
| 1 | 1 | 3.4 | 4 | 4.8 | 4.32 | 4.57 | 4.1 | 4.38 | 4.31 | 3.5 | 4.4 | 4.6168 | 11 | 6.26 | 5.16 | < 20 | 4.2 | < 10 | 5.1 | 3.8 | < 10 | 4.1 | < 5 |  |  | 1
| 1 | 2 | 3.5 | 3.9 | 5 | 4.35 | 4.62 | 4.04 | 4.28 | 4.23 | 3.4 | 4.25 | 4.6083 | 11 | 6.59 | 5.27 | < 20 | 4.2 | < 10 | 4.8 | 4.6 | < 10 | 4.2 | < 5 |  |  | 41
| 1 | 3 | 3.7 | 3.9 | 4.9 | 4.32 | 4.68 | 4.14 | 4.3 | 4.38 | 3.4 | 4.16 | 4.7123 | 12 | 7.1 | 5.28 | < 20 | 4.2 | < 10 | 5.1 | 3.9 | < 10 | 4.2 | < 5 |  |  | 16
| 1 | 4 | 3.7 | 4.2 | 5 | 4.33 | 4.79 | 4.12 | 4.33 | 4.35 | 3.3 | 4.31 | 4.5455 | 10 | 7.25 | 5.31 | < 20 | 4.1 | < 10 | 4.8 | 4.1 | < 10 | 4.2 | < 5 |  |  | 4.3141
| 1 | 5 | 3.5 | 4.4 | 5 | 4.32 | 4.67 | 4.16 | 4.21 | 4.19 | 3.3 | 4.34 | 4.6605 | 12 | 6.68 | 5.17 | < 20 | 4.1 | < 10 | 4.9 | 4 | < 10 | 4.2 | < 5 |  |  | 109
| 1 | 6 | 3.7 | 4.1 | 4.9 | 4.26 | 4.53 | 4.19 | 4.36 | 4.31 | 3.4 | 4.17 | 4.6773 | 10 | 6.78 | 5.07 | < 20 | 4.3 | < 10 | 5 | 3.8 | < 10 | 4.9 | < 5
| Mean |  | 3.5833 | 4.0833 | 4.9333 | 4.3167 | 4.6433 | 4.125 | 4.31 | 4.295 | 3.3833 | 4.2717 | 4.6368 | 11 | 6.7767 | 5.21 |  | 4.1833 |  | 4.95 | 4.0333 |  | 4.3
| Median |  | 3.6 | 4.05 | 4.95 | 4.32 | 4.645 | 4.13 | 4.315 | 4.31 | 3.4 | 4.28 | 4.6386 | 11 | 6.73 | 5.22 |  | 4.2 |  | 4.95 | 3.95 |  | 4.2
| Std Dev. |  | 0.13292 | 0.19408 | 0.08165 | 0.030111 | 0.092014 | 0.052058 | 0.061319 | 0.072042 | 0.075277 | 0.095795 | 0.059072 | 0.89443 | 0.3577 | 0.091869 |  | 0.075277 |  | 0.13784 | 0.30111 |  | 0.29665
| Rel.Std.Dev. |  | 3.709% | 4.753% | 1.655% | 0.6975% | 1.982% | 1.262% | 1.423% | 1.677% | 2.225% | 2.243% | 1.274% | 8.131% | 5.278% | 1.763% |  | 1.799% |  | 2.785% | 7.466% |  | 6.899%
| PDM3 |  | -16.94% | -5.348% | 14.35% | 0.06019% | 7.632% | -4.383% | -0.09434% | -0.442% | -21.57% | -0.9829% | 7.481% | 155% | 57.08% | 20.77% |  | -3.03% |  | 14.74% | -6.507% |  | -0.3261%
| Z-Score (Absolute) |  | 1.47 | 0.46 | 1.26 | 0.01 | 0.67 | 0.37 | 0 | 0.03 | 1.87 | 0.08 | 0.66 | NA | 4.99 | 1.82 | 11.51 | 0.26 | 1.4 | 1.29 | 0.56 | 1.4 | 0.02 | 3.66
Results from laboratory 13 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 611 (Certified Value 3.09 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g |  |  |  |  | 3
| 1 | 1 | 3.31 | 3.15 | 3.45 | 2.89 | 3.01 | 3.05 | 3.07 | 2.74 | 3.25 | 3.09 | 3.126 | 2.8 | 2.8369 | 3.1 | 3.33 | 3.11 | 2.91 | < 5 | 3.4 | 3 |  |  |  |  | 1
| 1 | 2 | 3.51 | 2.94 | 3.4 | 2.91 | 3.11 | 2.99 | 3.2 | 2.75 | 3.24 | 3.08 | 3.0343 | 2.79 | 2.7848 | 3.1 | 3.47 | 3.12 | 2.83 | < 5 | 3.4 | 3 |  |  |  |  | 42
| 1 | 3 | 3.57 | 3.19 | 3.35 | 2.91 | 3.12 | 3.01 | 2.99 | 2.75 | 3.22 | 3.07 | 3.021 | 2.74 | 2.7755 | 2.9 | 3.47 | 3.12 | 2.95 | < 5 | 3.4 | 3 |  |  |  |  | 16
| 1 | 4 | 3.56 | 3.22 | 3.25 | 2.98 | 3.15 | 3 | 3.04 | 2.66 | 3.27 | 3.11 | 2.949 | 2.85 | 2.83 | 3.1 | 3.44 | 3.03 | 2.82 | < 5 | 3.5 | 3 |  |  |  |  | 3.0946
| 1 | 5 | 3.46 | 3.45 | 3.28 | 2.98 | 3.15 | 3.15 | 2.88 | 2.61 | 3.26 | 3.06 | 3.1259 | 2.82 | 2.8426 | 2.9 | 3.51 | 2.99 | 2.91 | < 5 | 3.6 | 3 |  |  |  |  | 110
| 1 | 6 | 3.48 | 3.11 | 3.2 | 3.02 | 3.09 | 3.1 | 3.02 | 2.71 | 3.29 | 3 | 3.0693 | 2.8 | 2.6834 | 3.1 | 3.62 | 3.11 | 2.93 | < 5 | 3.2 | 3
| Mean |  | 3.4817 | 3.1767 | 3.3217 | 2.9483 | 3.105 | 3.05 | 3.0333 | 2.7033 | 3.255 | 3.0683 | 3.0542 | 2.8 | 2.7922 | 3.0333 | 3.4733 | 3.08 | 2.8917 |  | 3.4167 | 3
| Median |  | 3.495 | 3.17 | 3.315 | 2.945 | 3.115 | 3.03 | 3.03 | 2.725 | 3.255 | 3.075 | 3.0518 | 2.8 | 2.8074 | 3.1 | 3.47 | 3.11 | 2.91 |  | 3.4 | 3
| Std Dev. |  | 0.094534 | 0.16609 | 0.094956 | 0.051929 | 0.052058 | 0.063561 | 0.10463 | 0.057155 | 0.02429 | 0.037639 | 0.067958 | 0.036332 | 0.060236 | 0.10328 | 0.094375 | 0.055857 | 0.053821 |  | 0.13292 | 0
| Rel.Std.Dev. |  | 2.715% | 5.229% | 2.859% | 1.761% | 1.677% | 2.084% | 3.449% | 2.114% | 0.7462% | 1.227% | 2.225% | 1.298% | 2.157% | 3.405% | 2.717% | 1.814% | 1.861% |  | 3.89% | 0%
| PDM3 |  | 12.51% | 2.652% | 7.338% | -4.726% | 0.3361% | -1.441% | -1.98% | -12.64% | 5.183% | -0.8487% | -1.304% | -9.52% | -9.772% | -1.98% | 12.24% | -0.4717% | -6.558% |  | 10.41% | -3.057%
| Z-Score (Absolute) |  | 1.77 | 0.51 | 1.11 | 0.44 | 0.21 | 0.02 | 0.09 | 1.46 | 0.83 | 0.06 | 0 | 1.05 | 1.09 | 0.09 | 1.74 | 0.11 | 0.67 | 2.3 | 1.5 | NA
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 611 (Indicative Value 0.2 ppm) | [25 empty] | Indicative
Yb | Round | Replicate | 2.02: Lab; 2.03: Lab; 2.04: Lab; 2.05: Lab |  |  |  |  | 1
| No | No | 2.02: 2.04; 2.03: 2.18; 2.04: 2.20; 2.05: 2.21 |  |  |  |  | ppm
|  |  | 2.02: AR*MS; 2.03: AR*OES/MS; 2.04: AR*OES/MS; 2.05: AR*OES/MS |  |  |  |  | 2
|  |  | 2.02: 0.5g; 2.03: 0.2g; 2.04: 0.25g; 2.05: 0.25g |  |  |  |  | 2
| 1 | 1 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 1
| 1 | 2 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 13
| 1 | 3 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 16
| 1 | 4 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 0.2
| 1 | 5 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2 |  |  |  |  | 19
| 1 | 6 | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2
| Mean |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2
| Median |  | 2.02: 0.2; 2.03: 0.2; 2.04: 0.2; 2.05: 0.2
| Std Dev. |  | 2.02: 3.0405e-17; 2.03: 3.0405e-17; 2.04: 3.0405e-17; 2.05: 3.0405e-17
| Rel.Std.Dev. |  | 2.02: 1.52e-14%; 2.03: 1.52e-14%; 2.04: 1.52e-14%; 2.05: 1.52e-14%
| PDM3 |  | 2.02: -1.11e-14%; 2.03: -1.11e-14%; 2.04: -1.11e-14%; 2.05: -1.11e-14%
Z-Score (Absolute) | bar: 2.02 to 2.05
| Analytical results for Zn in OREAS 611 (Certified Value 2058 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.01 | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.13 | 2.14 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 | ppm
|  |  | AR*OES | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 0
|  |  | 0.25g | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g | 0
| 1 | 1 | 2180 | 2101 | 2184.4 | 1940 | 1980 | 2040 | 1960 | 2030 | 2030 | 2090 | 2032 | 1987.2029 | 2293 | 1956.7 | 2080 | 2090.7044 | 2018 | 2086.03 | 1998 | 2206 | 2010 | 2191 | 1960 | 2027 | 1
| 1 | 2 | 2110 | 2163 | 2148.8 | 2050 | 2000 | 2040 | 1960 | 2020 | 2070 | 2070 | 2069 | 1915.3686 | 2292 | 1956.4 | 2040 | 2083.7593 | 2039 | 2081.623 | 2050 | 2096 | 1950 | 2220 | 1960 | 2062 | 22
| 1 | 3 | 2070 | 2149 | 2159.7 | 1980 | 2000 | 2060 | 1970 | 2020 | 2070 | 2100 | 2054.8 | 1916.5343 | 2291 | 2013.8 | 2080 | 2109.5003 | 1980 | 2074.88 | 2010 | 2218 | 2070 | 2197 | 1960 | 2047 | 16
| 1 | 4 | 2150 | 2188 | 2180.5 | 1960 | 2000 | 2070 | 1960 | 2060 | 2050 | 2060 | 2105.1 | 1898.7542 | 2358 | 2039.6 | 2170 | 2107.3915 | 2043 | 2068.624 | 1980 | 2094 | 2070 | 2227 | 1970 | 2032 | 2057.7708
| 1 | 5 | 2190 | 2151 | 2189.7 | 1880 | 2000 | 2110 | 2000 | 2040 | 2060 | 2110 | 2039.5 | 1906.0725 | 2346 | 2002.4 | 2160 | 2102.3627 | 2006 | 2098.45 | 1995 | 2209 | 2050 | 2233 | 1970 | 2067 | 111
| 1 | 6 | 2120 | 2186 | 2177.1 | 1930 | 2030 | 2040 | 1990 | 1970 | 2020 | 2060 | 1988.4 | 1972.2286 | 2313 | 1997.1 | 2070 | 2040.2203 | 2032 | 2125.659 | 1922 | 2219 | 2080 | 2158 | 1960 | 2082
| Mean |  | 2136.6667 | 2156.3333 | 2173.3667 | 1956.6667 | 2001.6667 | 2060 | 1973.3333 | 2023.3333 | 2050 | 2081.6667 | 2048.1333 | 1932.6935 | 2315.5 | 1994.3333 | 2100 | 2088.9897 | 2019.6667 | 2089.211 | 1992.5 | 2173.6667 | 2038.3333 | 2204.3333 | 1963.3333 | 2052.8333
| Median |  | 2135 | 2157 | 2178.8 | 1950 | 2000 | 2050 | 1965 | 2025 | 2055 | 2080 | 2047.15 | 1915.9515 | 2303 | 1999.75 | 2080 | 2096.5335 | 2025 | 2083.8265 | 1996.5 | 2207.5 | 2060 | 2208.5 | 1960 | 2054.5
| Std Dev. |  | 45.4606 | 31.8476 | 15.7711 | 56.8038 | 16.0208 | 27.5681 | 17.5119 | 30.1109 | 20.9762 | 21.3698 | 39.0588 | 37.2991 | 29.6699 | 32.729 | 52.5357 | 25.8724 | 23.8048 | 20.5461 | 41.8939 | 61.1446 | 49.9667 | 28.1401 | 5.164 | 21.3112
| Rel.Std.Dev. |  | 2.128% | 1.477% | 0.7257% | 2.903% | 0.8004% | 1.338% | 0.8874% | 1.488% | 1.023% | 1.027% | 1.907% | 1.93% | 1.281% | 1.641% | 2.502% | 1.239% | 1.179% | 0.9834% | 2.103% | 2.813% | 2.451% | 1.277% | 0.263% | 1.038%
| PDM3 |  | 3.834% | 4.79% | 5.618% | -4.913% | -2.726% | 0.1083% | -4.103% | -1.674% | -0.3776% | 1.161% | -0.4683% | -6.078% | 12.52% | -3.083% | 2.052% | 1.517% | -1.852% | 1.528% | -3.172% | 5.632% | -0.9446% | 7.122% | -4.589% | -0.2399%
| Z-Score (Absolute) |  | 1.08 | 1.32 | 1.54 | 1.2 | 0.63 | 0.11 | 0.99 | 0.35 | 0.02 | 0.38 | 0.04 | 1.5 | 3.33 | 0.72 | 0.61 | 0.47 | 0.4 | 0.48 | 0.74 | 1.54 | 0.17 | 1.93 | 1.11 | 0.02
| Analytical results for Zr in OREAS 611 (Certified Value 12.5 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 2.02 | 2.03 | 2.04 | 2.05 | 2.06 | 2.07 | 2.08 | 2.09 | 2.10 | 2.11 | 2.12 | 2.15 | 2.17 | 2.18 | 2.19 | 2.20 | 2.21 | 2.22 | 2.23 | 2.24 | 2.25 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
|  |  | 01g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.2g | 0.2g | 0.25g | 0.25g | 0.25g | 10g | 0.25g | 0.5g | 01g |  |  |  | 2
| 1 | 1 | 13.2 | 12.2 | 17.7 | 11.8 | 12.7 | 13.2 | 13.2 | 11.9 | 11.5 | 14.03 | 10.9514 | 12.2 | 13.4581 | 11.2 | 11.89 | 11.7 | 11.3 | 12 | 20.9 | 11 | 14.3 |  |  |  | 1
| 1 | 2 | 13.9 | 12.9 | 17.6 | 12 | 13.1 | 13 | 14.9 | 12.1 | 11.4 | 14 | 11.0859 | 11.6 | 13.435 | 11.3 | 12.64 | 11.6 | 11.8 | 11.8 | 21 | 11 | 14.71 |  |  |  | 23
| 1 | 3 | 14.3 | 12.7 | 17.2 | 12.1 | 13.3 | 13.3 | 13.5 | 11.8 | 11.3 | 13.86 | 11.0828 | 11.5 | 13.4943 | 10.9 | 12.93 | 11.8 | 11.8 | 12.5 | 21.2 | 11 | 14.12 |  |  |  | 16
| 1 | 4 | 14.3 | 12.9 | 16.8 | 12.4 | 13.5 | 13.1 | 13.2 | 11.6 | 11.3 | 14.06 | 10.9645 | 11.6 | 13.4555 | 11.4 | 12.77 | 11.2 | 11.8 | 12.4 | 21.5 | 11 | 14.53 |  |  |  | 12.5377
| 1 | 5 | 13.7 | 13.5 | 16.2 | 12.5 | 13.3 | 13.4 | 13.1 | 11.4 | 11.5 | 13.85 | 11.0951 | 11.7 | 13.6895 | 10.9 | 12.83 | 11.1 | 11.7 | 12.3 | 22.3 | 11 | 14.94 |  |  |  | 112
| 1 | 6 | 13.9 | 12.6 | 16.3 | 12.4 | 13.4 | 13.2 | 12.8 | 11.8 | 11.6 | 13.6 | 11.0408 | 11.7 | 13.1918 | 11.1 | 12.76 | 11.8 | 11.9 | 12.5 | 15.8 | 11 | 14.35
| Mean |  | 13.8833 | 12.8 | 16.9667 | 12.2 | 13.2167 | 13.2 | 13.45 | 11.7667 | 11.4333 | 13.9 | 11.0368 | 11.7167 | 13.454 | 11.1333 | 12.6367 | 11.5333 | 11.7167 | 12.25 | 20.45 | 11 | 14.4917
| Median |  | 13.9 | 12.8 | 17 | 12.25 | 13.3 | 13.2 | 13.2 | 11.8 | 11.45 | 13.93 | 11.0618 | 11.65 | 13.4568 | 11.15 | 12.765 | 11.65 | 11.8 | 12.35 | 21.1 | 11 | 14.44
| Std Dev. |  | 0.41191 | 0.42895 | 0.64083 | 0.27568 | 0.28577 | 0.14142 | 0.74498 | 0.24221 | 0.12111 | 0.17123 | 0.063967 | 0.24833 | 0.15887 | 0.20656 | 0.37787 | 0.30768 | 0.2137 | 0.2881 | 2.333 | 0 | 0.29836
| Rel.Std.Dev. |  | 2.967% | 3.351% | 3.777% | 2.26% | 2.162% | 1.071% | 5.539% | 2.058% | 1.059% | 1.232% | 0.5796% | 2.119% | 1.181% | 1.855% | 2.99% | 2.668% | 1.824% | 2.352% | 11.41% | 0% | 2.059%
| PDM3 |  | 10.73% | 2.092% | 35.33% | -2.693% | 5.415% | 5.283% | 7.277% | -6.15% | -8.808% | 10.87% | -11.97% | -6.548% | 7.309% | -11.2% | 0.7894% | -8.011% | -6.548% | -2.295% | 63.11% | -12.26% | 15.58%
| Z-Score (Absolute) |  | 0.78 | 0.05 | 2.86 | 0.35 | 0.34 | 0.32 | 0.49 | 0.64 | 0.87 | 0.8 | 1.13 | 0.67 | 0.5 | 1.07 | 0.05 | 0.8 | 0.67 | 0.32 | 5.2 | NA | 1.19
Results from laboratory 24 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1059: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1077: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1095: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1151: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 611 (Indicative Value 11.28 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.24 | 1
| 1 | 2 | 11.32 | 16
| Mean |  | 11.28 | 16
| Median |  | 11.28 | 11.28
| Std Dev. |  | 0.056569 | 22
| Rel.Std.Dev. |  | 0.5015%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 611 (Indicative Value 3635 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 3640 | 1
| 1 | 2 | 3630 | 17
| Mean |  | 3635 | 16
| Median |  | 3635 | 3635
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 0.1945%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for BaO in OREAS 611 (Indicative Value 5075 ppm) |  |  | Indicative
BaO | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 5130 | 1
| 1 | 2 | 5020 | 18
| Mean |  | 5075 | 16
| Median |  | 5075 | 5075
| Std Dev. |  | 77.7817 | 24
| Rel.Std.Dev. |  | 1.533%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 611 (Indicative Value 0.31 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.31 | 1
| 1 | 2 | 0.31 | 19
| Mean |  | 0.31 | 16
| Median |  | 0.31 | 0.31
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 611 (Indicative Value 30 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 40 | 20
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 14.1421 | 26
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 611 (Indicative Value 15 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 10 | 21
| Mean |  | 15 | 16
| Median |  | 15 | 15
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr2O3 in OREAS 611 (Indicative Value 90 ppm) |  |  | Indicative
Cr2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 100 | 22
| Mean |  | 90 | 16
| Median |  | 90 | 90
| Std Dev. |  | 14.1421 | 28
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 611 (Indicative Value 1.17 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.17 | 1
| 1 | 2 | 1.16 | 23
| Mean |  | 1.165 | 16
| Median |  | 1.165 | 1.165
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 0.607%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 611 (Indicative Value 3.62 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.62 | 1
| 1 | 2 | 3.62 | 16
| Mean |  | 3.62 | 16
| Median |  | 3.62 | 3.62
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 611 (Indicative Value 2.29 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.29 | 1
| 1 | 2 | 2.29 | 17
| Mean |  | 2.29 | 16
| Median |  | 2.29 | 2.29
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 611 (Indicative Value 0.275 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.27 | 1
| 1 | 2 | 0.28 | 18
| Mean |  | 0.275 | 16
| Median |  | 0.275 | 0.275
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 2.571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 611 (Indicative Value 0.012 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.013 | 1
| 1 | 2 | 0.011 | 19
| Mean |  | 0.012 | 16
| Median |  | 0.012 | 0.012
| Std Dev. |  | 0.0014142 | 25
| Rel.Std.Dev. |  | 11.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 611 (Indicative Value 1.13 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.12 | 1
| 1 | 2 | 1.14 | 20
| Mean |  | 1.13 | 16
| Median |  | 1.13 | 1.13
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 611 (Indicative Value 35 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 40 | 21
| Mean |  | 35 | 16
| Median |  | 35 | 35
| Std Dev. |  | 7.0711 | 27
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 611 (Indicative Value 0.129 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.128 | 1
| 1 | 2 | 0.129 | 22
| Mean |  | 0.1285 | 16
| Median |  | 0.1285 | 0.1285
| Std Dev. |  | 0.00070711 | 28
| Rel.Std.Dev. |  | 0.5503%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 611 (Indicative Value 690 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 680 | 1
| 1 | 2 | 700 | 23
| Mean |  | 690 | 16
| Median |  | 690 | 690
| Std Dev. |  | 14.1421 | 29
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 611 (Indicative Value 4.33 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 4.3288 | 1
| 1 | 2 | 4.3248 | 16
| Mean |  | 4.3268 | 16
| Median |  | 4.3268 | 4.3268
| Std Dev. |  | 0.0028284 | 22
| Rel.Std.Dev. |  | 0.06537%
| PDM3 |  | -0.001065%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 611 (Indicative Value 69.54 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 69.48 | 1
| 1 | 2 | 69.59 | 17
| Mean |  | 69.535 | 16
| Median |  | 69.535 | 69.535
| Std Dev. |  | 0.077782 | 23
| Rel.Std.Dev. |  | 0.1119%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 611 (Indicative Value 20 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 18
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 611 (Indicative Value 338 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 340 | 1
| 1 | 2 | 340 | 19
| Mean |  | 340 | 16
| Median |  | 340 | 338.2378
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.521%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 611 (Indicative Value 0.323 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.322 | 1
| 1 | 2 | 0.324 | 20
| Mean |  | 0.323 | 16
| Median |  | 0.323 | 0.323
| Std Dev. |  | 0.0014142 | 26
| Rel.Std.Dev. |  | 0.4378%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 611 (Indicative Value 55 ppm) |  |  | Indicative
V2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 40 | 1
| 1 | 2 | 70 | 21
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 21.2132 | 27
| Rel.Std.Dev. |  | 38.57%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 611 (Indicative Value 2080 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 2080 | 1
| 1 | 2 | 2080 | 22
| Mean |  | 2080 | 16
| Median |  | 2080 | 2080
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 611 (Indicative Value 196 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 192 | 1
| 1 | 2 | 200 | 23
| Mean |  | 196 | 16
| Median |  | 196 | 196.1853
| Std Dev. |  | 5.6569 | 29
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | -0.09447%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 611 (Indicative Value 8.48 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 8.46 | 1
| 1 | 2 | 8.5 | 25
| Mean |  | 8.48 | 16
| Median |  | 8.48 | 8.48
| Std Dev. |  | 0.028284 | 31
| Rel.Std.Dev. |  | 0.3335%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Cu in OREAS 611 (Indicative Value 1.15 wt.%) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 2.22 | wt.%
|  |  | PF*OES | 2
2
| 1 | 1 | 1.14 | 1
| 1 | 2 | 1.14 | 23
| 1 | 3 | 1.14 | 16
| 1 | 4 | 1.15 | 1.1483
| 1 | 5 | 1.13 | 33
| 1 | 6 | 1.19
| Mean |  | 1.1483
| Median |  | 1.14
| Std Dev. |  | 0.02137
| Rel.Std.Dev. |  | 1.861%
| PDM3 |  | 2.887e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 611 (Indicative Value 96.3 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 93 | 1
| 1 | 2 | 99.5 | 29
| Mean |  | 96.25 | 16
| Median |  | 96.25 | 96.25
| Std Dev. |  | 4.5962 | 35
| Rel.Std.Dev. |  | 4.775%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 611 (Indicative Value 3550 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 3530 | 1
| 1 | 2 | 3570 | 17
| Mean |  | 3550 | 16
| Median |  | 3550 | 3550
| Std Dev. |  | 28.2843 | 36
| Rel.Std.Dev. |  | 0.7967%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 611 (Indicative Value 4445 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 4410 | 1
| 1 | 2 | 4480 | 5
| Mean |  | 4445 | 16
| Median |  | 4445 | 4445
| Std Dev. |  | 49.4975 | 37
| Rel.Std.Dev. |  | 1.114%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 611 (Indicative Value 1.7 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.8 | 32
| Mean |  | 1.7 | 16
| Median |  | 1.7 | 1.7
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 611 (Indicative Value 260 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 261 | 1
| 1 | 2 | 258 | 33
| Mean |  | 259.5 | 16
| Median |  | 259.5 | 259.5
| Std Dev. |  | 2.1213 | 39
| Rel.Std.Dev. |  | 0.8175%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 611 (Indicative Value 16.8 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.9 | 1
| 1 | 2 | 17.6 | 34
| Mean |  | 16.75 | 16
| Median |  | 16.75 | 16.75
| Std Dev. |  | 1.2021 | 40
| Rel.Std.Dev. |  | 7.177%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 611 (Indicative Value 51 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 51.3 | 1
| 1 | 2 | 49.7 | 35
| Mean |  | 50.5 | 16
| Median |  | 50.5 | 50.5
| Std Dev. |  | 1.1314 | 41
| Rel.Std.Dev. |  | 2.24%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 611 (Indicative Value 8.55 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.8 | 1
| 1 | 2 | 8.3 | 21
| Mean |  | 8.55 | 16
| Median |  | 8.55 | 8.55
| Std Dev. |  | 0.35355 | 42
| Rel.Std.Dev. |  | 4.135%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 611 (Indicative Value 61 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 64 | 1
| 1 | 2 | 57 | 37
| Mean |  | 60.5 | 16
| Median |  | 60.5 | 60.5
| Std Dev. |  | 4.9497 | 43
| Rel.Std.Dev. |  | 8.181%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 611 (Indicative Value 2.01 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.99 | 1
| 1 | 2 | 2.02 | 38
| Mean |  | 2.005 | 16
| Median |  | 2.005 | 2.005
| Std Dev. |  | 0.021213 | 44
| Rel.Std.Dev. |  | 1.058%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 611 (Indicative Value 1.17 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.15 | 1
| 1 | 2 | 1.18 | 23
| Mean |  | 1.165 | 16
| Median |  | 1.165 | 1.165
| Std Dev. |  | 0.021213 | 45
| Rel.Std.Dev. |  | 1.821%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 611 (Indicative Value 1.66 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.68 | 1
| 1 | 2 | 1.64 | 10
| Mean |  | 1.66 | 16
| Median |  | 1.66 | 1.66
| Std Dev. |  | 0.028284 | 46
| Rel.Std.Dev. |  | 1.704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 611 (Indicative Value 0.72 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.69 | 1
| 1 | 2 | 0.74 | 11
| Mean |  | 0.715 | 16
| Median |  | 0.715 | 0.715
| Std Dev. |  | 0.035355 | 47
| Rel.Std.Dev. |  | 4.945%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 611 (Indicative Value 0.79 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.78 | 1
| 1 | 2 | 0.8 | 12
| Mean |  | 0.79 | 16
| Median |  | 0.79 | 0.79
| Std Dev. |  | 0.014142 | 48
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 611 (Indicative Value 23.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23.1 | 1
| 1 | 2 | 23.5 | 43
| Mean |  | 23.3 | 16
| Median |  | 23.3 | 23.3
| Std Dev. |  | 0.28284 | 49
| Rel.Std.Dev. |  | 1.214%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 611 (Indicative Value 3.02 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.92 | 1
| 1 | 2 | 3.11 | 13
| Mean |  | 3.015 | 16
| Median |  | 3.015 | 3.015
| Std Dev. |  | 0.13435 | 50
| Rel.Std.Dev. |  | 4.456%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 611 (Indicative Value 5.13 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.1 | 1
| 1 | 2 | 5.15 | 6
| Mean |  | 5.125 | 16
| Median |  | 5.125 | 5.125
| Std Dev. |  | 0.035355 | 51
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 611 (Indicative Value 5.22 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.27 | 1
| 1 | 2 | 5.16 | 29
| Mean |  | 5.215 | 16
| Median |  | 5.215 | 5.215
| Std Dev. |  | 0.077782 | 35
| Rel.Std.Dev. |  | 1.492%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 611 (Indicative Value 0.26 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.24 | 1
| 1 | 2 | 0.27 | 8
| Mean |  | 0.255 | 16
| Median |  | 0.255 | 0.255
| Std Dev. |  | 0.021213 | 36
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 611 (Indicative Value 4.3 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.3 | 1
| 1 | 2 | 4.3 | 31
| Mean |  | 4.3 | 16
| Median |  | 4.3 | 4.3
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 611 (Indicative Value 26.6 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.9 | 1
| 1 | 2 | 26.2 | 32
| Mean |  | 26.55 | 16
| Median |  | 26.55 | 26.55
| Std Dev. |  | 0.49497 | 38
| Rel.Std.Dev. |  | 1.864%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 611 (Indicative Value 105 ppb) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 110 | 1
| 1 | 2 | 100 | 6
| Mean |  | 105 | 16
| Median |  | 105 | 105
| Std Dev. |  | 7.0711 | 39
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 611 (Indicative Value 75 ppm) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 75 | 1
| 1 | 2 | 75 | 34
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 611 (Indicative Value 5.2 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.4 | 1
| 1 | 2 | 5 | 35
| Mean |  | 5.2 | 16
| Median |  | 5.2 | 5.2
| Std Dev. |  | 0.28284 | 41
| Rel.Std.Dev. |  | 5.439%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 611 (Indicative Value 9.24 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.28 | 1
| 1 | 2 | 9.19 | 36
| Mean |  | 9.235 | 16
| Median |  | 9.235 | 9.235
| Std Dev. |  | 0.06364 | 42
| Rel.Std.Dev. |  | 0.6891%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 611 (Indicative Value 21.1 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21.5 | 1
| 1 | 2 | 20.7 | 10
| Mean |  | 21.1 | 16
| Median |  | 21.1 | 21.1
| Std Dev. |  | 0.56569 | 43
| Rel.Std.Dev. |  | 2.681%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 611 (Indicative Value 32 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32 | 1
| 1 | 2 | 32 | 21
| Mean |  | 32 | 16
| Median |  | 32 | 32
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 611 (Indicative Value 699 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 699 | 1
| 1 | 2 | 698 | 23
| Mean |  | 698.5 | 16
| Median |  | 698.5 | 698.5
| Std Dev. |  | 0.70711 | 45
| Rel.Std.Dev. |  | 0.1012%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 611 (Indicative Value 5.69 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.77 | 1
| 1 | 2 | 5.61 | 12
| Mean |  | 5.69 | 16
| Median |  | 5.69 | 5.69
| Std Dev. |  | 0.11314 | 46
| Rel.Std.Dev. |  | 1.988%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 611 (Indicative Value 61 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 60.8 | 1
| 1 | 2 | 61.4 | 41
| Mean |  | 61.1 | 16
| Median |  | 61.1 | 61.1
| Std Dev. |  | 0.42426 | 47
| Rel.Std.Dev. |  | 0.6944%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 611 (Indicative Value 17.5 ppb) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | < 10 | 14
| Mean |  | 30 | 16
| Median |  | 30 | 17.5
| Std Dev. |  |  | 48
Rel.Std.Dev.
| PDM3 |  | 71.43%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 611 (Indicative Value 387 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 385 | 1
| 1 | 2 | 388 | 43
| Mean |  | 386.5 | 16
| Median |  | 386.5 | 386.5
| Std Dev. |  | 2.1213 | 49
| Rel.Std.Dev. |  | 0.5489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 611 (Indicative Value 3.75 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 3.7 | 44
| Mean |  | 3.75 | 16
| Median |  | 3.75 | 3.75
| Std Dev. |  | 0.070711 | 50
| Rel.Std.Dev. |  | 1.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 611 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 45
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 51
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 611 (Indicative Value 4.09 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.14 | 1
| 1 | 2 | 4.04 | 8
| Mean |  | 4.09 | 16
| Median |  | 4.09 | 4.09
| Std Dev. |  | 0.070711 | 35
| Rel.Std.Dev. |  | 1.729%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 611 (Indicative Value 32.1 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.8 | 1
| 1 | 2 | 32.4 | 18
| Mean |  | 32.1 | 16
| Median |  | 32.1 | 32.1
| Std Dev. |  | 0.42426 | 36
| Rel.Std.Dev. |  | 1.322%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 611 (Indicative Value 349 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 351 | 1
| 1 | 2 | 347 | 19
| Mean |  | 349 | 16
| Median |  | 349 | 349
| Std Dev. |  | 2.8284 | 37
| Rel.Std.Dev. |  | 0.8104%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 611 (Indicative Value 0.77 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.77 | 1
| 1 | 2 | 0.77 | 9
| Mean |  | 0.77 | 16
| Median |  | 0.77 | 0.77
| Std Dev. |  | 0 | 38
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 611 (Indicative Value 0.38 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.38 | 10
| Mean |  | 0.38 | 16
| Median |  | 0.38 | 0.38
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 611 (Indicative Value 53 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 52.4 | 1
| 1 | 2 | 53.6 | 34
| Mean |  | 53 | 16
| Median |  | 53 | 53
| Std Dev. |  | 0.84853 | 40
| Rel.Std.Dev. |  | 1.601%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 611 (Indicative Value 10.4 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.6 | 1
| 1 | 2 | 10.1 | 35
| Mean |  | 10.35 | 16
| Median |  | 10.35 | 10.35
| Std Dev. |  | 0.35355 | 41
| Rel.Std.Dev. |  | 3.416%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 611 (Indicative Value 0.192 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.194 | 1
| 1 | 2 | 0.19 | 11
| Mean |  | 0.192 | 16
| Median |  | 0.192 | 0.192
| Std Dev. |  | 0.0028284 | 42
| Rel.Std.Dev. |  | 1.473%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 611 (Indicative Value 2.9 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3 | 37
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.14142 | 43
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 611 (Indicative Value 100 ppb) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppb
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 6
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 611 (Indicative Value 2.76 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.76 | 1
| 1 | 2 | 2.75 | 39
| Mean |  | 2.755 | 16
| Median |  | 2.755 | 2.755
| Std Dev. |  | 0.0070711 | 45
| Rel.Std.Dev. |  | 0.2567%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 611 (Indicative Value 33 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.2 | 1
| 1 | 2 | 32.8 | 40
| Mean |  | 33 | 16
| Median |  | 33 | 33
| Std Dev. |  | 0.28284 | 46
| Rel.Std.Dev. |  | 0.8571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 611 (Indicative Value 8.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9 | 1
| 1 | 2 | 8.5 | 41
| Mean |  | 8.75 | 16
| Median |  | 8.75 | 8.75
| Std Dev. |  | 0.35355 | 47
| Rel.Std.Dev. |  | 4.041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 611 (Indicative Value 7.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.62 | 1
| 1 | 2 | 7.58 | 42
| Mean |  | 7.6 | 16
| Median |  | 7.6 | 7.6
| Std Dev. |  | 0.028284 | 48
| Rel.Std.Dev. |  | 0.3722%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 611 (Indicative Value 0.72 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.74 | 1
| 1 | 2 | 0.7 | 13
| Mean |  | 0.72 | 16
| Median |  | 0.72 | 0.72
| Std Dev. |  | 0.028284 | 49
| Rel.Std.Dev. |  | 3.928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 611 (Indicative Value 2015 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2010 | 1
| 1 | 2 | 2020 | 22
| Mean |  | 2015 | 16
| Median |  | 2015 | 2015
| Std Dev. |  | 7.0711 | 50
| Rel.Std.Dev. |  | 0.3509%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 611 (Indicative Value 179 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 178 | 1
| 1 | 2 | 179 | 23
| Mean |  | 178.5 | 16
| Median |  | 178.5 | 178.5
| Std Dev. |  | 0.70711 | 51
| Rel.Std.Dev. |  | 0.3961%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## XRPA
| Analytical results for Au in OREAS 611 (Indicative Value 15.81 ppm) |  |  |  | Indicative
Au | Round | Replicate | Chrysos | Chrysos | 1
| No | No | Perth | Kalgoorlie | ppm
|  |  | Raw*XRPA | Raw*XRPA | 2
2
| 1 | 1 | 15.907 | 15.709 | 1
| Mean |  | 15.907 | 15.709 | 47
| Median |  | 15.907 | 15.709 | 16
| Std Dev. |  |  |  | 15.808
| Rel.Std.Dev. |  |  |  | 53
| PDM3 |  | 0.6263% | -0.6263%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C12: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.