## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 554
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Al2O3, wt.% | 8.2301 | 0.090524 | 8.049 | 8.4111 | 7.9585 | 8.5016 | 1.1% | 2.2% | 3.3% | 7.8186 | 8.6416
BaO, ppm | 1417.9097 | 68.834 | 1280.2416 | 1555.5778 | 1211.4076 | 1624.4119 | 4.855% | 9.709% | 14.56% | 1347.0143 | 1488.8052
CaO, wt.% | 5.5801 | 0.074242 | 5.4316 | 5.7285 | 5.3573 | 5.8028 | 1.33% | 2.661% | 3.991% | 5.3011 | 5.8591
Co, wt.% | 0.80452 | 0.020807 | 0.76291 | 0.84613 | 0.7421 | 0.86694 | 2.586% | 5.173% | 7.759% | 0.76429 | 0.84475
Cr2O3, ppm | 97.1022 | 18.0837 | 60.9349 | 133.2696 | 42.8512 | 151.3533 | 18.62% | 37.25% | 55.87% | 92.2471 | 101.9573
Cu, wt.% | 1.5786 | 0.027709 | 1.5231 | 1.634 | 1.4954 | 1.6617 | 1.755% | 3.511% | 5.266% | 1.4996 | 1.6575
Fe2O3, wt.% | 2.5749 | 0.035292 | 2.5044 | 2.6455 | 2.4691 | 2.6808 | 1.371% | 2.741% | 4.112% | 2.4462 | 2.7037
K2O, wt.% | 2.0056 | 0.026904 | 1.9518 | 2.0594 | 1.9249 | 2.0863 | 1.341% | 2.683% | 4.024% | 1.9053 | 2.1059
MgO, wt.% | 5.2726 | 0.061269 | 5.1501 | 5.3951 | 5.0888 | 5.4564 | 1.162% | 2.324% | 3.486% | 5.009 | 5.5362
MnO, wt.% | 0.14408 | 0.0061317 | 0.13182 | 0.15635 | 0.12569 | 0.16248 | 4.256% | 8.511% | 12.77% | 0.13688 | 0.15129
P2O5, wt.% | 0.07316 | 0.0066246 | 0.059911 | 0.086409 | 0.053286 | 0.093034 | 9.055% | 18.11% | 27.17% | 0.069502 | 0.076818
SiO2, wt.% | 60.2623 | 0.48012 | 59.302 | 61.2225 | 58.8219 | 61.7026 | 0.7967% | 1.593% | 2.39% | 57.2492 | 63.2754
SO3, wt.% | 1.9838 | 0.031049 | 1.9217 | 2.0459 | 1.8906 | 2.0769 | 1.565% | 3.13% | 4.695% | 1.8846 | 2.083
SrO, ppm | 94.0241 | 13.2881 | 67.4479 | 120.6002 | 54.1598 | 133.8883 | 14.13% | 28.27% | 42.4% | 89.3229 | 98.7253
TiO2, wt.% | 0.3818 | 0.011494 | 0.35881 | 0.40479 | 0.34732 | 0.41628 | 3.01% | 6.021% | 9.031% | 0.36271 | 0.40089
Thermogravimetry
LOI1000, wt.% | 10.1141 | 0.25144 | 9.6112 | 10.617 | 9.3598 | 10.8684 | 2.486% | 4.972% | 7.458% | 9.6084 | 10.6198
Infrared Combustion
C, wt.% | 2.8436 | 0.05586 | 2.7318 | 2.9553 | 2.676 | 3.0111 | 1.964% | 3.929% | 5.893% | 2.7014 | 2.9857
S, wt.% | 0.77127 | 0.020473 | 0.73033 | 0.81222 | 0.70985 | 0.83269 | 2.654% | 5.309% | 7.963% | 0.73271 | 0.80983
Sulphuric Acid 5% Leach
Co, wt.% | 0.55549 | 0.042462 | 0.47056 | 0.64041 | 0.4281 | 0.68287 | 7.644% | 15.29% | 22.93% | 0.52771 | 0.58326
Sulphuric Acid 10% Leach
Co, wt.% | 0.54576 | 0.041644 | 0.46248 | 0.62905 | 0.42083 | 0.6707 | 7.63% | 15.26% | 22.89% | 0.51848 | 0.57305
Peroxide Fusion ICP
Ag, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Al, wt.% | 4.3246 | 0.13988 | 4.0448 | 4.6043 | 3.9049 | 4.7442 | 3.234% | 6.469% | 9.703% | 4.1083 | 4.5408
As, ppm | 65.869 | 5.1807 | 55.5076 | 76.2305 | 50.3269 | 81.4112 | 7.865% | 15.73% | 23.6% | 62.5756 | 69.1625
B, ppm | 178.3333 | 19.5551 | 139.2232 | 217.4434 | 119.6682 | 236.9985 | 10.97% | 21.93% | 32.9% | 169.4167 | 187.25
Ba, ppm | 1280.486 | 67.6266 | 1145.2329 | 1415.7391 | 1077.6063 | 1483.3657 | 5.281% | 10.56% | 15.84% | 1216.4617 | 1344.5103
Bi, ppm | 2.8264 | 0.52931 | 1.7677 | 3.885 | 1.2384 | 4.4143 | 18.73% | 37.46% | 56.18% | 2.685 | 2.9677
Ca, wt.% | 3.9556 | 0.16561 | 3.6244 | 4.2868 | 3.4588 | 4.4524 | 4.187% | 8.373% | 12.56% | 3.7578 | 4.1534
Ce, ppm | 67.7987 | 2.2731 | 63.2526 | 72.3449 | 60.9795 | 74.618 | 3.353% | 6.705% | 10.06% | 64.4088 | 71.1887
Co, wt.% | 0.80949 | 0.024217 | 0.76105 | 0.85792 | 0.73684 | 0.88214 | 2.992% | 5.983% | 8.975% | 0.76901 | 0.84996
Cr, ppm | 68.6766 | 13.6435 | 41.3897 | 95.9635 | 27.7463 | 109.607 | 19.87% | 39.73% | 59.6% | 65.2428 | 72.1105
Cs, ppm | 2.5548 | 0.22546 | 2.1038 | 3.0057 | 1.8784 | 3.2311 | 8.825% | 17.65% | 26.47% | 2.427 | 2.6825
Cu, wt.% | 1.5693 | 0.042559 | 1.4841 | 1.6544 | 1.4416 | 1.6969 | 2.712% | 5.424% | 8.136% | 1.4908 | 1.6477
Dy, ppm | 3.5352 | 0.17614 | 3.1829 | 3.8875 | 3.0068 | 4.0636 | 4.982% | 9.965% | 14.95% | 3.3584 | 3.712
Er, ppm | 1.9881 | 0.11702 | 1.7541 | 2.2222 | 1.6371 | 2.3392 | 5.886% | 11.77% | 17.66% | 1.8887 | 2.0875
Eu, ppm | 0.84767 | 0.13492 | 0.57782 | 1.1175 | 0.44289 | 1.2524 | 15.92% | 31.83% | 47.75% | 0.80528 | 0.89005
Fe, wt.% | 1.8192 | 0.04822 | 1.7228 | 1.9156 | 1.6745 | 1.9639 | 2.651% | 5.301% | 7.952% | 1.7282 | 1.9102
Ga, ppm | 13.5184 | 1.5557 | 10.4071 | 16.6298 | 8.8515 | 18.1854 | 11.51% | 23.02% | 34.52% | 12.8425 | 14.1944
Gd, ppm | 4.2692 | 0.43498 | 3.3993 | 5.1392 | 2.9643 | 5.5742 | 10.19% | 20.38% | 30.57% | 4.0558 | 4.4827
Hf, ppm | 2.24 | 0.42149 | 1.397 | 3.083 | 0.97553 | 3.5045 | 18.82% | 37.63% | 56.45% | 2.128 | 2.352
Ho, ppm | 0.72467 | 0.049531 | 0.6256 | 0.82373 | 0.57607 | 0.87326 | 6.835% | 13.67% | 20.51% | 0.68843 | 0.7609
K, wt.% | 1.6921 | 0.063156 | 1.5658 | 1.8184 | 1.5026 | 1.8815 | 3.732% | 7.465% | 11.2% | 1.6075 | 1.7767
La, ppm | 35.1863 | 1.9868 | 31.2127 | 39.1599 | 29.2258 | 41.1468 | 5.647% | 11.29% | 16.94% | 33.427 | 36.9456
Li, ppm | 34.2376 | 6.4904 | 21.2567 | 47.2184 | 14.7663 | 53.7088 | 18.96% | 37.91% | 56.87% | 32.5257 | 35.9495
Lu, ppm | 0.2945 | 0.034066 | 0.22637 | 0.36263 | 0.1923 | 0.3967 | 11.57% | 23.13% | 34.7% | 0.27977 | 0.30923
Mg, wt.% | 3.1429 | 0.062264 | 3.0184 | 3.2675 | 2.9561 | 3.3297 | 1.981% | 3.962% | 5.943% | 2.9858 | 3.3001
Mn, wt.% | 0.1103 | 0.0050361 | 0.10023 | 0.12037 | 0.095189 | 0.12541 | 4.566% | 9.132% | 13.7% | 0.10478 | 0.11581
Nd, ppm | 28.1822 | 1.3037 | 25.5748 | 30.7897 | 24.271 | 32.0934 | 4.626% | 9.252% | 13.88% | 26.7731 | 29.5913
Ni, ppm | 33.6704 | 8.1352 | 17.3999 | 49.9409 | 9.2647 | 58.0761 | 24.16% | 48.32% | 72.48% | 31.9869 | 35.3539
P, wt.% | 0.033883 | 0.0048622 | 0.024159 | 0.043608 | 0.019297 | 0.04847 | 14.35% | 28.7% | 43.05% | 0.032189 | 0.035577
Pb, ppm | 18.6149 | 3.017 | 12.5808 | 24.6489 | 9.5638 | 27.6659 | 16.21% | 32.42% | 48.62% | 17.6841 | 19.5456
Pr, ppm | 7.8567 | 0.48936 | 6.8779 | 8.8354 | 6.3886 | 9.3247 | 6.229% | 12.46% | 18.69% | 7.4638 | 8.2495
Rb, ppm | 88.8619 | 2.6134 | 83.6351 | 94.0887 | 81.0217 | 96.7022 | 2.941% | 5.882% | 8.823% | 84.4188 | 93.305
S, wt.% | 0.76948 | 0.019611 | 0.73026 | 0.80871 | 0.71065 | 0.82832 | 2.549% | 5.097% | 7.646% | 0.73101 | 0.80796
Sb, ppm | 2.2418 | 0.43279 | 1.3762 | 3.1074 | 0.94343 | 3.5401 | 19.31% | 38.61% | 57.92% | 2.1297 | 2.3539
Sc, ppm | 7.9771 | 1.0699 | 5.8374 | 10.1169 | 4.7675 | 11.1868 | 13.41% | 26.82% | 40.24% | 7.5783 | 8.376
Si, wt.% | 28.7481 | 0.63527 | 27.4775 | 30.0186 | 26.8423 | 30.6539 | 2.21% | 4.42% | 6.629% | 27.3107 | 30.1855
Sm, ppm | 4.9298 | 0.39488 | 4.14 | 5.7196 | 3.7452 | 6.1145 | 8.01% | 16.02% | 24.03% | 4.6833 | 5.1763
Sr, ppm | 71.0197 | 3.0863 | 64.8472 | 77.1922 | 61.7609 | 80.2785 | 4.346% | 8.691% | 13.04% | 67.4687 | 74.5707
Ta, ppm | 0.69694 | 0.1559 | 0.38514 | 1.0087 | 0.22924 | 1.1647 | 22.37% | 44.74% | 67.11% | 0.6621 | 0.73179
Tb, ppm | 0.59194 | 0.096987 | 0.39797 | 0.78592 | 0.30098 | 0.8829 | 16.38% | 32.77% | 49.15% | 0.56235 | 0.62154
Th, ppm | 9.2941 | 0.67048 | 7.9531 | 10.635 | 7.2826 | 11.3055 | 7.214% | 14.43% | 21.64% | 8.8294 | 9.7588
Ti, wt.% | 0.22793 | 0.010429 | 0.20707 | 0.24879 | 0.19664 | 0.25922 | 4.575% | 9.151% | 13.73% | 0.21653 | 0.23933
Tm, ppm | 0.30033 | 0.024842 | 0.25065 | 0.35002 | 0.22581 | 0.37486 | 8.271% | 16.54% | 24.81% | 0.28532 | 0.31535
U, ppm | 14.2717 | 0.6415 | 12.9887 | 15.5547 | 12.3472 | 16.1962 | 4.495% | 8.99% | 13.48% | 13.5581 | 14.9853
V, ppm | 115.9134 | 6.74 | 102.4334 | 129.3934 | 95.6934 | 136.1333 | 5.815% | 11.63% | 17.44% | 110.1177 | 121.709
Y, ppm | 20.1249 | 1.1277 | 17.8695 | 22.3802 | 16.7418 | 23.5079 | 5.603% | 11.21% | 16.81% | 19.1186 | 21.1311
Yb, ppm | 1.949 | 0.19626 | 1.5564 | 2.3415 | 1.3602 | 2.5377 | 10.07% | 20.14% | 30.21% | 1.8515 | 2.0464
Zn, ppm | 56.0093 | 10.9357 | 34.1378 | 77.8807 | 23.202 | 88.8165 | 19.52% | 39.05% | 58.57% | 53.2088 | 58.8097
Zr, ppm | 81.1687 | 10.8919 | 59.3849 | 102.9525 | 48.4929 | 113.8444 | 13.42% | 26.84% | 40.26% | 77.1102 | 85.2271
4-Acid Digestion
Ag, ppm | 0.25471 | 0.023901 | 0.20691 | 0.30251 | 0.18301 | 0.32641 | 9.384% | 18.77% | 28.15% | 0.24197 | 0.26744
Al, wt.% | 4.3167 | 0.13644 | 4.0438 | 4.5896 | 3.9074 | 4.726 | 3.161% | 6.322% | 9.482% | 4.1009 | 4.5326
As, ppm | 65.9925 | 2.807 | 60.3784 | 71.6066 | 57.5713 | 74.4136 | 4.254% | 8.507% | 12.76% | 62.6928 | 69.2921
Ba, ppm | 1274.3611 | 49.6024 | 1175.1562 | 1373.5659 | 1125.5538 | 1423.1683 | 3.892% | 7.785% | 11.68% | 1210.643 | 1338.0791
Be, ppm | 1.8399 | 0.12364 | 1.5926 | 2.0872 | 1.469 | 2.2108 | 6.72% | 13.44% | 20.16% | 1.7479 | 1.9319
Bi, ppm | 2.7354 | 0.18317 | 2.3691 | 3.1017 | 2.1859 | 3.2849 | 6.696% | 13.39% | 20.09% | 2.5986 | 2.8722
Ca, wt.% | 3.96 | 0.1133 | 3.7334 | 4.1866 | 3.6201 | 4.2999 | 2.861% | 5.722% | 8.583% | 3.762 | 4.158
Cd, ppm | 0.20734 | 0.020469 | 0.1664 | 0.24827 | 0.14593 | 0.26874 | 9.873% | 19.75% | 29.62% | 0.19697 | 0.2177
Ce, ppm | 63.7732 | 3.9265 | 55.9202 | 71.6262 | 51.9937 | 75.5527 | 6.157% | 12.31% | 18.47% | 60.5845 | 66.9618
Co, wt.% | 0.81545 | 0.03355 | 0.74835 | 0.88255 | 0.7148 | 0.91611 | 4.114% | 8.229% | 12.34% | 0.77468 | 0.85623
Cr, ppm | 49.4667 | 5.5952 | 38.2763 | 60.657 | 32.6812 | 66.2522 | 11.31% | 22.62% | 33.93% | 46.9933 | 51.94
Cs, ppm | 2.4249 | 0.11258 | 2.1997 | 2.6501 | 2.0872 | 2.7626 | 4.643% | 9.285% | 13.93% | 2.3037 | 2.5461
Cu, wt.% | 1.5881 | 0.024344 | 1.5394 | 1.6368 | 1.5151 | 1.6612 | 1.533% | 3.066% | 4.599% | 1.5087 | 1.6675
Dy, ppm | 2.7091 | 0.28382 | 2.1415 | 3.2768 | 1.8577 | 3.5606 | 10.48% | 20.95% | 31.43% | 2.5737 | 2.8446
Er, ppm | 1.4316 | 0.21781 | 0.99595 | 1.8672 | 0.77814 | 2.085 | 15.21% | 30.43% | 45.64% | 1.36 | 1.5031
Eu, ppm | 0.8243 | 0.032236 | 0.75983 | 0.88877 | 0.72759 | 0.92101 | 3.911% | 7.821% | 11.73% | 0.78309 | 0.86552
Fe, wt.% | 1.8036 | 0.068451 | 1.6667 | 1.9405 | 1.5983 | 2.009 | 3.795% | 7.59% | 11.39% | 1.7134 | 1.8938
Ga, ppm | 12.5157 | 0.60836 | 11.2989 | 13.7324 | 10.6906 | 14.3407 | 4.861% | 9.722% | 14.58% | 11.8899 | 13.1414
Gd, ppm | 3.7817 | 0.19875 | 3.3842 | 4.1791 | 3.1854 | 4.3779 | 5.256% | 10.51% | 15.77% | 3.5926 | 3.9707
Ge, ppm | 0.094857 | 0.020509 | 0.05384 | 0.13587 | 0.033331 | 0.15638 | 21.62% | 43.24% | 64.86% | 0.090114 | 0.0996
Hf, ppm | 1.877 | 0.19844 | 1.4801 | 2.2739 | 1.2817 | 2.4723 | 10.57% | 21.14% | 31.72% | 1.7831 | 1.9708
Ho, ppm | 0.52471 | 0.075975 | 0.37276 | 0.67666 | 0.29679 | 0.75264 | 14.48% | 28.96% | 43.44% | 0.49848 | 0.55095
In, ppm | 0.10494 | 0.0065249 | 0.09189 | 0.11799 | 0.085365 | 0.12451 | 6.218% | 12.44% | 18.65% | 0.099693 | 0.11019
K, wt.% | 1.7074 | 0.065116 | 1.5772 | 1.8377 | 1.5121 | 1.9028 | 3.814% | 7.627% | 11.44% | 1.6221 | 1.7928
La, ppm | 30.5763 | 2.4147 | 25.7469 | 35.4057 | 23.3322 | 37.8204 | 7.897% | 15.79% | 23.69% | 29.0475 | 32.1051
Li, ppm | 33.87 | 1.683 | 30.5041 | 37.2359 | 28.8211 | 38.9189 | 4.969% | 9.938% | 14.91% | 32.1765 | 35.5635
Lu, ppm | 0.21143 | 0.015233 | 0.18096 | 0.2419 | 0.16573 | 0.25713 | 7.205% | 14.41% | 21.61% | 0.20086 | 0.222
Mg, wt.% | 3.1442 | 0.096346 | 2.9516 | 3.3369 | 2.8552 | 3.4333 | 3.064% | 6.128% | 9.193% | 2.987 | 3.3015
Mn, wt.% | 0.10817 | 0.004101 | 0.099971 | 0.11637 | 0.09587 | 0.12048 | 3.791% | 7.582% | 11.37% | 0.10276 | 0.11358
Mo, ppm | 3.5386 | 0.22355 | 3.0915 | 3.9857 | 2.868 | 4.2093 | 6.317% | 12.63% | 18.95% | 3.3617 | 3.7156
Na, wt.% | 0.059742 | 0.0067199 | 0.046303 | 0.073182 | 0.039583 | 0.079902 | 11.25% | 22.5% | 33.74% | 0.056755 | 0.062729
Nb, ppm | 5.0774 | 1.8045 | 1.4685 | 8.6864 | 0 | 10.4909 | 35.54% | 71.08% | 106.6% | 4.8236 | 5.3313
Nd, ppm | 25.299 | 1.9124 | 21.4741 | 29.1239 | 19.5617 | 31.0363 | 7.559% | 15.12% | 22.68% | 24.0341 | 26.564
Ni, ppm | 31.2769 | 1.3803 | 28.5163 | 34.0374 | 27.1361 | 35.4177 | 4.413% | 8.826% | 13.24% | 29.713 | 32.8407
P, wt.% | 0.030592 | 0.00097977 | 0.028632 | 0.032551 | 0.027652 | 0.033531 | 3.203% | 6.405% | 9.608% | 0.029062 | 0.032121
Pb, ppm | 16.5952 | 0.9198 | 14.7556 | 18.4347 | 13.8358 | 19.3545 | 5.543% | 11.09% | 16.63% | 15.7654 | 17.4249
Pr, ppm | 6.8615 | 0.41284 | 6.0358 | 7.6872 | 5.6229 | 8.1 | 6.017% | 12.03% | 18.05% | 6.5184 | 7.2045
Rb, ppm | 87.9223 | 3.9457 | 80.031 | 95.8137 | 76.0853 | 99.7594 | 4.488% | 8.975% | 13.46% | 83.5262 | 92.3185
Re, ppm | 0.043952 | 0.0032022 | 0.037548 | 0.050357 | 0.034346 | 0.053559 | 7.286% | 14.57% | 21.86% | 0.041755 | 0.04615
S, wt.% | 0.77837 | 0.022541 | 0.73328 | 0.82345 | 0.71074 | 0.84599 | 2.896% | 5.792% | 8.688% | 0.73945 | 0.81728
Sb, ppm | 2.2204 | 0.19709 | 1.8263 | 2.6146 | 1.6292 | 2.8117 | 8.876% | 17.75% | 26.63% | 2.1094 | 2.3315
Sc, ppm | 8.1008 | 0.21939 | 7.662 | 8.5395 | 7.4426 | 8.7589 | 2.708% | 5.416% | 8.125% | 7.6957 | 8.5058
Se, ppm | 5.2926 | 0.75262 | 3.7873 | 6.7978 | 3.0347 | 7.5504 | 14.22% | 28.44% | 42.66% | 5.0279 | 5.5572
Sm, ppm | 4.8186 | 0.20021 | 4.4181 | 5.219 | 4.2179 | 5.4192 | 4.155% | 8.31% | 12.47% | 4.5776 | 5.0595
Sn, ppm | 2.2152 | 0.15815 | 1.8989 | 2.5315 | 1.7407 | 2.6896 | 7.139% | 14.28% | 21.42% | 2.1044 | 2.3259
Sr, ppm | 67.204 | 2.7187 | 61.7666 | 72.6414 | 59.048 | 75.3601 | 4.045% | 8.091% | 12.14% | 63.8438 | 70.5642
Ta, ppm | 0.34833 | 0.10488 | 0.13857 | 0.55808 | 0.033696 | 0.66296 | 30.11% | 60.22% | 90.33% | 0.33091 | 0.36574
Tb, ppm | 0.51813 | 0.040117 | 0.43789 | 0.59836 | 0.39777 | 0.63848 | 7.743% | 15.49% | 23.23% | 0.49222 | 0.54403
Te, ppm | 0.049133 | 0.019119 | 0.010896 | 0.08737 | 0 | 0.10649 | 38.91% | 77.82% | 116.7% | 0.046677 | 0.05159
Th, ppm | 8.9941 | 0.62238 | 7.7493 | 10.2389 | 7.127 | 10.8612 | 6.92% | 13.84% | 20.76% | 8.5444 | 9.4438
Ti, wt.% | 0.1609 | 0.034763 | 0.09137 | 0.23042 | 0.056608 | 0.26518 | 21.61% | 43.21% | 64.82% | 0.15285 | 0.16894
Tl, ppm | 0.52737 | 0.030327 | 0.46671 | 0.58802 | 0.43639 | 0.61835 | 5.751% | 11.5% | 17.25% | 0.501 | 0.55373
Tm, ppm | 0.20233 | 0.015687 | 0.17096 | 0.23371 | 0.15527 | 0.2494 | 7.753% | 15.51% | 23.26% | 0.19222 | 0.21245
U, ppm | 13.7747 | 0.97679 | 11.8211 | 15.7283 | 10.8443 | 16.7051 | 7.091% | 14.18% | 21.27% | 13.086 | 14.4635
V, ppm | 112.2119 | 4.4439 | 103.3242 | 121.0996 | 98.8804 | 125.5435 | 3.96% | 7.92% | 11.88% | 106.6013 | 117.8225
W, ppm | 1.2697 | 0.18331 | 0.90313 | 1.6364 | 0.71983 | 1.8197 | 14.44% | 28.87% | 43.31% | 1.2063 | 1.3332
Y, ppm | 12.9139 | 0.92534 | 11.0632 | 14.7646 | 10.1379 | 15.6899 | 7.165% | 14.33% | 21.5% | 12.2682 | 13.5596
Yb, ppm | 1.4162 | 0.19589 | 1.0245 | 1.808 | 0.82857 | 2.0039 | 13.83% | 27.66% | 41.5% | 1.3454 | 1.4871
Zn, ppm | 55.2939 | 2.9423 | 49.4093 | 61.1784 | 46.467 | 64.1207 | 5.321% | 10.64% | 15.96% | 52.5292 | 58.0586
Zr, ppm | 62.518 | 3.1097 | 56.2987 | 68.7373 | 53.189 | 71.847 | 4.974% | 9.948% | 14.92% | 59.3921 | 65.6439
Aqua Regia Digestion
Ag, ppm | 0.17772 | 0.011877 | 0.15396 | 0.20147 | 0.14209 | 0.21335 | 6.683% | 13.37% | 20.05% | 0.16883 | 0.1866
Al, wt.% | 0.5945 | 0.045659 | 0.50318 | 0.68582 | 0.45752 | 0.73148 | 7.68% | 15.36% | 23.04% | 0.56477 | 0.62422
As, ppm | 64.8996 | 3.2837 | 58.3321 | 71.4671 | 55.0484 | 74.7509 | 5.06% | 10.12% | 15.18% | 61.6546 | 68.1446
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 225.9637 | 22.7901 | 180.3836 | 271.5438 | 157.5935 | 294.3339 | 10.09% | 20.17% | 30.26% | 214.6655 | 237.2619
Be, ppm | 0.62889 | 0.043751 | 0.54139 | 0.71639 | 0.49764 | 0.76014 | 6.957% | 13.91% | 20.87% | 0.59744 | 0.66033
Bi, ppm | 2.5797 | 0.11727 | 2.3451 | 2.8142 | 2.2279 | 2.9315 | 4.546% | 9.092% | 13.64% | 2.4507 | 2.7087
Ca, wt.% | 3.8542 | 0.12368 | 3.6068 | 4.1015 | 3.4832 | 4.2252 | 3.209% | 6.418% | 9.627% | 3.6615 | 4.0469
Cd, ppm | 0.19873 | 0.017261 | 0.16421 | 0.23326 | 0.14695 | 0.25052 | 8.686% | 17.37% | 26.06% | 0.1888 | 0.20867
Ce, ppm | 34.7548 | 3.8554 | 27.044 | 42.4655 | 23.1886 | 46.3209 | 11.09% | 22.19% | 33.28% | 33.017 | 36.4925
Co, wt.% | 0.76093 | 0.042072 | 0.67679 | 0.84508 | 0.63472 | 0.88715 | 5.529% | 11.06% | 16.59% | 0.72289 | 0.79898
Cr, ppm | 15.5232 | 1.9088 | 11.7057 | 19.3407 | 9.7969 | 21.2495 | 12.3% | 24.59% | 36.89% | 14.747 | 16.2993
Cs, ppm | 0.56649 | 0.0995 | 0.36749 | 0.76549 | 0.26799 | 0.86499 | 17.56% | 35.13% | 52.69% | 0.53817 | 0.59482
Cu, wt.% | 1.5785 | 0.053706 | 1.4711 | 1.686 | 1.4174 | 1.7397 | 3.402% | 6.804% | 10.21% | 1.4996 | 1.6575
Dy, ppm | 1.343 | 0.14334 | 1.0563 | 1.6296 | 0.91295 | 1.773 | 10.67% | 21.35% | 32.02% | 1.2758 | 1.4101
Er, ppm | 0.57173 | 0.038502 | 0.49473 | 0.64874 | 0.45623 | 0.68724 | 6.734% | 13.47% | 20.2% | 0.54315 | 0.60032
Eu, ppm | 0.4545 | 0.081269 | 0.29196 | 0.61704 | 0.21069 | 0.69831 | 17.88% | 35.76% | 53.64% | 0.43178 | 0.47723
Fe, wt.% | 1.6391 | 0.051919 | 1.5353 | 1.743 | 1.4834 | 1.7949 | 3.167% | 6.335% | 9.502% | 1.5572 | 1.7211
Ga, ppm | 1.7642 | 0.26589 | 1.2324 | 2.296 | 0.96651 | 2.5619 | 15.07% | 30.14% | 45.22% | 1.676 | 1.8524
Gd, ppm | 2.1755 | 0.44542 | 1.2847 | 3.0664 | 0.83927 | 3.5118 | 20.47% | 40.95% | 61.42% | 2.0668 | 2.2843
Hf, ppm | 0.26654 | 0.030401 | 0.20574 | 0.32734 | 0.17534 | 0.35774 | 11.41% | 22.81% | 34.22% | 0.25321 | 0.27987
Hg, ppm | 0.018636 | 0.0058438 | 0.0069488 | 0.030324 | 0.001105 | 0.036168 | 31.36% | 62.71% | 94.07% | 0.017705 | 0.019568
Ho, ppm | 0.22235 | 0.027902 | 0.16654 | 0.27815 | 0.13864 | 0.30606 | 12.55% | 25.1% | 37.65% | 0.21123 | 0.23347
In, ppm | 0.063143 | 0.0037112 | 0.055721 | 0.070566 | 0.05201 | 0.074277 | 5.877% | 11.75% | 17.63% | 0.059986 | 0.0663
K, wt.% | 0.17697 | 0.018372 | 0.14023 | 0.21371 | 0.12185 | 0.23209 | 10.38% | 20.76% | 31.14% | 0.16812 | 0.18582
La, ppm | 14.2363 | 1.9387 | 10.3588 | 18.1138 | 8.4201 | 20.0525 | 13.62% | 27.24% | 40.85% | 13.5245 | 14.9481
Li, ppm | 18.3578 | 2.5644 | 13.229 | 23.4867 | 10.6645 | 26.0511 | 13.97% | 27.94% | 41.91% | 17.44 | 19.2757
Lu, ppm | 0.079333 | 0.011211 | 0.056911 | 0.10176 | 0.0457 | 0.11297 | 14.13% | 28.26% | 42.39% | 0.075367 | 0.0833
Mg, wt.% | 2.9015 | 0.09994 | 2.7016 | 3.1013 | 2.6016 | 3.2013 | 3.444% | 6.889% | 10.33% | 2.7564 | 3.0465
Mn, wt.% | 0.10631 | 0.0048488 | 0.096609 | 0.116 | 0.09176 | 0.12085 | 4.561% | 9.122% | 13.68% | 0.10099 | 0.11162
Mo, ppm | 3.3198 | 0.1703 | 2.9792 | 3.6604 | 2.8089 | 3.8307 | 5.13% | 10.26% | 15.39% | 3.1538 | 3.4857
Nd, ppm | 12.9641 | 1.7092 | 9.5457 | 16.3825 | 7.8365 | 18.0917 | 13.18% | 26.37% | 39.55% | 12.3159 | 13.6123
Ni, ppm | 28.388 | 2.2648 | 23.8583 | 32.9177 | 21.5935 | 35.1826 | 7.978% | 15.96% | 23.93% | 26.9686 | 29.8074
P, wt.% | 0.026448 | 0.0012862 | 0.023876 | 0.029021 | 0.02259 | 0.030307 | 4.863% | 9.726% | 14.59% | 0.025126 | 0.027771
Pb, ppm | 10.9759 | 1.3443 | 8.2873 | 13.6644 | 6.9431 | 15.0087 | 12.25% | 24.5% | 36.74% | 10.4271 | 11.5247
Pr, ppm | 2.8174 | 0.82315 | 1.1711 | 4.4637 | 0.3479 | 5.2868 | 29.22% | 58.43% | 87.65% | 2.6765 | 2.9582
Rb, ppm | 8.8308 | 0.87296 | 7.0848 | 10.5767 | 6.2119 | 11.4497 | 9.885% | 19.77% | 29.66% | 8.3892 | 9.2723
Re, ppm | 0.044583 | 0.0025335 | 0.039516 | 0.04965 | 0.036983 | 0.052184 | 5.683% | 11.37% | 17.05% | 0.042354 | 0.046812
S, wt.% | 0.74281 | 0.048901 | 0.64501 | 0.84061 | 0.59611 | 0.88951 | 6.583% | 13.17% | 19.75% | 0.70567 | 0.77995
Sb, ppm | 1.4854 | 0.18792 | 1.1096 | 1.8613 | 0.92166 | 2.0492 | 12.65% | 25.3% | 37.95% | 1.4112 | 1.5597
Sc, ppm | 2.2326 | 0.16329 | 1.906 | 2.5592 | 1.7427 | 2.7225 | 7.314% | 14.63% | 21.94% | 2.121 | 2.3442
Se, ppm | 4.8965 | 0.47817 | 3.9401 | 5.8528 | 3.462 | 6.331 | 9.766% | 19.53% | 29.3% | 4.6517 | 5.1413
Sn, ppm | 0.31905 | 0.051731 | 0.21559 | 0.42251 | 0.16386 | 0.47424 | 16.21% | 32.43% | 48.64% | 0.3031 | 0.335
Sr, ppm | 40.2326 | 2.3546 | 35.5235 | 44.9417 | 33.169 | 47.2963 | 5.852% | 11.7% | 17.56% | 38.221 | 42.2443
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.29164 | 0.022389 | 0.24686 | 0.33642 | 0.22447 | 0.35881 | 7.677% | 15.35% | 23.03% | 0.27706 | 0.30622
Te, ppm | 0.048241 | 0.01074 | 0.026761 | 0.069722 | 0.016021 | 0.080462 | 22.26% | 44.53% | 66.79% | 0.045829 | 0.050653
Th, ppm | 5.5111 | 0.25799 | 4.9951 | 6.0271 | 4.7371 | 6.2851 | 4.681% | 9.363% | 14.04% | 5.2355 | 5.7866
Tl, ppm | 0.10886 | 0.013126 | 0.082614 | 0.13512 | 0.069488 | 0.14824 | 12.06% | 24.11% | 36.17% | 0.10342 | 0.11431
U, ppm | 11.8427 | 0.84889 | 10.1449 | 13.5405 | 9.296 | 14.3893 | 7.168% | 14.34% | 21.5% | 11.2505 | 12.4348
V, ppm | 19.2956 | 1.4857 | 16.3242 | 22.267 | 14.8385 | 23.7527 | 7.7% | 15.4% | 23.1% | 18.3308 | 20.2604
W, ppm | 0.10471 | 0.011955 | 0.080798 | 0.12862 | 0.068843 | 0.14057 | 11.42% | 22.83% | 34.25% | 0.099473 | 0.10994
Y, ppm | 5.6926 | 0.37107 | 4.9505 | 6.4348 | 4.5794 | 6.8058 | 6.518% | 13.04% | 19.56% | 5.408 | 5.9773
Yb, ppm | 0.49776 | 0.0077751 | 0.48221 | 0.51331 | 0.47443 | 0.52108 | 1.562% | 3.124% | 4.686% | 0.47287 | 0.52264
Zn, ppm | 48.2215 | 3.2512 | 41.719 | 54.7239 | 38.4678 | 57.9752 | 6.742% | 13.48% | 20.23% | 45.8104 | 50.6325
Zr, ppm | 8.9917 | 0.39495 | 8.2018 | 9.7816 | 7.8068 | 10.1765 | 4.392% | 8.785% | 13.18% | 8.5421 | 9.4413

## Confidence & Tolerance Limits
Table 2. Certified Values, 95% Confidence and Tolerance Limits for OREAS 554
Constituent | Certified | 95% Confidence Limits |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 8.2301 | 8.19 | 8.2701 | 8.1481 | 8.3121
BaO, Barium oxide (ppm) | 1417.9097 | 1378.1599 | 1457.6596 | 1358.6966 | 1477.1229
CaO, Calcium oxide (wt.%) | 5.5801 | 5.5438 | 5.6164 | 5.5479 | 5.6122
Co, Cobalt (wt.%) | 0.80452 | 0.79256 | 0.81648 | 0.79244 | 0.81659
Cr2O3, Chromium(III) oxide (ppm) | 97.1022 | 86.0146 | 108.1899 | IND | IND
Cu, Copper (wt.%) | 1.5786 | 1.5604 | 1.5967 | 1.5581 | 1.599
Fe2O3, Iron(III) oxide (wt.%) | 2.5749 | 2.5579 | 2.592 | 2.5522 | 2.5977
K2O, Potassium oxide (wt.%) | 2.0056 | 1.9932 | 2.0181 | 1.9924 | 2.0188
MgO, Magnesium oxide (wt.%) | 5.2726 | 5.2431 | 5.3021 | 5.2354 | 5.3098
MnO, Manganese oxide (wt.%) | 0.14408 | 0.14137 | 0.14679 | 0.14178 | 0.14638
P2O5, Phosphorus(V) oxide (wt.%) | 0.07316 | 0.069771 | 0.076549 | 0.070725 | 0.075595
SiO2, Silicon dioxide (wt.%) | 60.2623 | 60.0781 | 60.4465 | 59.96 | 60.5645
SO3, Sulphur trioxide (wt.%) | 1.9838 | 1.9602 | 2.0073 | 1.9526 | 2.015
SrO, Strontium oxide (ppm) | 94.0241 | 84.2725 | 103.7756 | IND | IND
TiO2, Titanium dioxide (wt.%) | 0.3818 | 0.37708 | 0.38652 | 0.37161 | 0.39199
Thermogravimetry
LOI1000, Loss On Ignition @1000°C (wt.%) | 10.1141 | 10.0074 | 10.2208 | 10.0039 | 10.2243
Infrared Combustion
C, Carbon (wt.%) | 2.8436 | 2.8226 | 2.8645 | 2.8122 | 2.875
S, Sulphur (wt.%) | 0.77127 | 0.76338 | 0.77916 | 0.75504 | 0.78751
Sulphuric Acid 5% Leach
Co, Cobalt (wt.%) | 0.55549 | 0.51547 | 0.5955 | 0.53399 | 0.57698
Sulphuric Acid 10% Leach
Co, Cobalt (wt.%) | 0.54576 | 0.51907 | 0.57246 | 0.52856 | 0.56297
Peroxide Fusion ICP
Ag, Silver (ppm) | < 1 | IND | IND | IND | IND
Al, Aluminium (wt.%) | 4.3246 | 4.2623 | 4.3869 | 4.2411 | 4.4081
As, Arsenic (ppm) | 65.869 | 60.9 | 70.8381 | 63.1296 | 68.6085
B, Boron (ppm) | 178.3333 | 160.2609 | 196.4058 | IND | IND
Ba, Barium (ppm) | 1280.486 | 1245.6696 | 1315.3024 | 1252.2295 | 1308.7425
Bi, Bismuth (ppm) | 2.8264 | 2.4777 | 3.1751 | 2.5582 | 3.0946
Ca, Calcium (wt.%) | 3.9556 | 3.8806 | 4.0306 | 3.8753 | 4.0359
Ce, Cerium (ppm) | 67.7987 | 65.9792 | 69.6183 | 65.8354 | 69.7621
Co, Cobalt (wt.%) | 0.80949 | 0.79925 | 0.81972 | 0.79289 | 0.82609
Cr, Chromium (ppm) | 68.6766 | 60.0941 | 77.2592 | 62.6882 | 74.6651
Cs, Caesium (ppm) | 2.5548 | 2.492 | 2.6175 | 2.3503 | 2.7593
Cu, Copper (wt.%) | 1.5693 | 1.5505 | 1.588 | 1.5304 | 1.6081
Dy, Dysprosium (ppm) | 3.5352 | 3.4526 | 3.6179 | 3.3375 | 3.7329
Er, Erbium (ppm) | 1.9881 | 1.911 | 2.0652 | 1.851 | 2.1253
Eu, Europium (ppm) | 0.84767 | 0.67387 | 1.0215 | 0.78879 | 0.90655
Fe, Iron (wt.%) | 1.8192 | 1.7985 | 1.8399 | 1.7828 | 1.8556
Ga, Gallium (ppm) | 13.5184 | 12.2949 | 14.742 | 12.7931 | 14.2438
Gd, Gadolinium (ppm) | 4.2692 | 3.9106 | 4.6279 | 3.8981 | 4.6404
Hf, Hafnium (ppm) | 2.24 | 1.7695 | 2.7105 | IND | IND
Ho, Holmium (ppm) | 0.72467 | 0.69799 | 0.75135 | 0.67172 | 0.77762
K, Potassium (wt.%) | 1.6921 | 1.6623 | 1.7218 | 1.6352 | 1.7489
La, Lanthanum (ppm) | 35.1863 | 33.8089 | 36.5637 | 34.0264 | 36.3461
Li, Lithium (ppm) | 34.2376 | 29.8726 | 38.6026 | 32.6974 | 35.7778
Lu, Lutetium (ppm) | 0.2945 | 0.25826 | 0.33074 | 0.27484 | 0.31416
Mg, Magnesium (wt.%) | 3.1429 | 3.1182 | 3.1677 | 3.0958 | 3.1901
Mn, Manganese (wt.%) | 0.1103 | 0.10816 | 0.11244 | 0.10635 | 0.11424
Nd, Neodymium (ppm) | 28.1822 | 27.3342 | 29.0302 | 27.1275 | 29.2369
Ni, Nickel (ppm) | 33.6704 | 26.5289 | 40.8119 | 29.2419 | 38.0989
P, Phosphorus (wt.%) | 0.033883 | 0.029155 | 0.038611 | IND | IND
Pb, Lead (ppm) | 18.6149 | 16.614 | 20.6157 | 15.7531 | 21.4766
Pr, Praseodymium (ppm) | 7.8567 | 7.5416 | 8.1717 | 7.5886 | 8.1247
Rb, Rubidium (ppm) | 88.8619 | 87.6087 | 90.1152 | 85.9486 | 91.7753
S, Sulphur (wt.%) | 0.76948 | 0.75721 | 0.78175 | 0.75378 | 0.78518
Sb, Antimony (ppm) | 2.2418 | 1.6696 | 2.814 | 1.9827 | 2.5009
Sc, Scandium (ppm) | 7.9771 | 6.7881 | 9.1662 | IND | IND
Si, Silicon (wt.%) | 28.7481 | 28.3822 | 29.114 | 28.3683 | 29.1279
Sm, Samarium (ppm) | 4.9298 | 4.6751 | 5.1845 | 4.6637 | 5.196
Sr, Strontium (ppm) | 71.0197 | 69.1465 | 72.8929 | 68.5714 | 73.468
Ta, Tantalum (ppm) | 0.69694 | 0.52832 | 0.86557 | IND | IND
Tb, Terbium (ppm) | 0.59194 | 0.48626 | 0.69763 | 0.55312 | 0.63077
Th, Thorium (ppm) | 9.2941 | 8.9092 | 9.679 | 8.7715 | 9.8167
Ti, Titanium (wt.%) | 0.22793 | 0.22339 | 0.23246 | 0.22029 | 0.23557
Tm, Thulium (ppm) | 0.30033 | 0.27217 | 0.3285 | 0.27629 | 0.32438
U, Uranium (ppm) | 14.2717 | 13.8009 | 14.7425 | 13.6303 | 14.9131
V, Vanadium (ppm) | 115.9134 | 112.567 | 119.2598 | 110.6125 | 121.2142
Y, Yttrium (ppm) | 20.1249 | 19.5133 | 20.7365 | 19.237 | 21.0127
Yb, Ytterbium (ppm) | 1.949 | 1.824 | 2.0739 | 1.7138 | 2.1841
Zn, Zinc (ppm) | 56.0093 | 49.4307 | 62.5878 | 48.9032 | 63.1153
Zr, Zirconium (ppm) | 81.1687 | 67.2139 | 95.1234 | 76.3177 | 86.0197
4-Acid Digestion
Ag, Silver (ppm) | 0.25471 | 0.24188 | 0.26753 | 0.23113 | 0.27829
Al, Aluminium (wt.%) | 4.3167 | 4.2646 | 4.3689 | 4.2373 | 4.3961
As, Arsenic (ppm) | 65.9925 | 64.863 | 67.1219 | 64.1661 | 67.8188
Ba, Barium (ppm) | 1274.3611 | 1251.5171 | 1297.2051 | 1249.7671 | 1298.955
Be, Beryllium (ppm) | 1.8399 | 1.785 | 1.8948 | 1.7964 | 1.8834
Bi, Bismuth (ppm) | 2.7354 | 2.6516 | 2.8192 | 2.6449 | 2.8259
Ca, Calcium (wt.%) | 3.96 | 3.9193 | 4.0008 | 3.8841 | 4.036
Cd, Cadmium (ppm) | 0.20734 | 0.20022 | 0.21445 | 0.17768 | 0.23699
Ce, Cerium (ppm) | 63.7732 | 61.9304 | 65.616 | 62.1436 | 65.4027
Co, Cobalt (wt.%) | 0.81545 | 0.80191 | 0.829 | 0.79559 | 0.83532
Cr, Chromium (ppm) | 49.4667 | 47.2666 | 51.6667 | 45.2879 | 53.6455
Cs, Caesium (ppm) | 2.4249 | 2.3747 | 2.4751 | 2.3595 | 2.4903
Cu, Copper (wt.%) | 1.5881 | 1.5793 | 1.5969 | 1.566 | 1.6103
Dy, Dysprosium (ppm) | 2.7091 | 2.512 | 2.9063 | 2.6045 | 2.8137
Er, Erbium (ppm) | 1.4316 | 1.2587 | 1.6044 | 1.3701 | 1.493
Eu, Europium (ppm) | 0.8243 | 0.78753 | 0.86107 | 0.79235 | 0.85625
Fe, Iron (wt.%) | 1.8036 | 1.775 | 1.8322 | 1.7695 | 1.8377
Ga, Gallium (ppm) | 12.5157 | 12.2634 | 12.7679 | 12.2047 | 12.8266
Gd, Gadolinium (ppm) | 3.7817 | 3.6291 | 3.9342 | 3.6147 | 3.9486
Ge, Germanium (ppm) | 0.094857 | 0.080052 | 0.10966 | IND | IND
Hf, Hafnium (ppm) | 1.877 | 1.793 | 1.961 | 1.7953 | 1.9587
Ho, Holmium (ppm) | 0.52471 | 0.4647 | 0.58472 | 0.49533 | 0.55409
In, Indium (ppm) | 0.10494 | 0.10267 | 0.10721 | 0.096592 | 0.11329
K, Potassium (wt.%) | 1.7074 | 1.6794 | 1.7355 | 1.6768 | 1.738
La, Lanthanum (ppm) | 30.5763 | 29.4517 | 31.7009 | 29.9285 | 31.2241
Li, Lithium (ppm) | 33.87 | 33.1429 | 34.597 | 32.9727 | 34.7672
Lu, Lutetium (ppm) | 0.21143 | 0.19792 | 0.22494 | 0.18745 | 0.2354
Mg, Magnesium (wt.%) | 3.1442 | 3.108 | 3.1805 | 3.0791 | 3.2093
Mn, Manganese (wt.%) | 0.10817 | 0.10655 | 0.10979 | 0.10557 | 0.11077
Mo, Molybdenum (ppm) | 3.5386 | 3.4466 | 3.6307 | 3.4106 | 3.6667
Na, Sodium (wt.%) | 0.059742 | 0.056459 | 0.063025 | 0.057341 | 0.062144
Nb, Niobium (ppm) | 5.0774 | 4.2756 | 5.8793 | 4.8614 | 5.2934
Nd, Neodymium (ppm) | 25.299 | 23.9983 | 26.5997 | 24.6166 | 25.9814
Ni, Nickel (ppm) | 31.2769 | 30.7358 | 31.818 | 30.1758 | 32.378
P, Phosphorus (wt.%) | 0.030592 | 0.030223 | 0.03096 | 0.029352 | 0.031832
Pb, Lead (ppm) | 16.5952 | 16.2234 | 16.9669 | 16.1147 | 17.0756
Pr, Praseodymium (ppm) | 6.8615 | 6.5674 | 7.1555 | 6.6436 | 7.0794
Rb, Rubidium (ppm) | 87.9223 | 86.3224 | 89.5223 | 86.1635 | 89.6812
Re, Rhenium (ppm) | 0.043952 | 0.04374 | 0.044165 | 0.039643 | 0.048262
S, Sulphur (wt.%) | 0.77837 | 0.76972 | 0.78701 | 0.75866 | 0.79807
Sb, Antimony (ppm) | 2.2204 | 2.1348 | 2.3061 | 2.1481 | 2.2928
Sc, Scandium (ppm) | 8.1008 | 7.9922 | 8.2093 | 7.8933 | 8.3082
Se, Selenium (ppm) | 5.2926 | 5.0338 | 5.5514 | 4.6947 | 5.8904
Sm, Samarium (ppm) | 4.8186 | 4.6876 | 4.9495 | 4.6199 | 5.0172
Sn, Tin (ppm) | 2.2152 | 2.1422 | 2.2882 | 2.1109 | 2.3194
Sr, Strontium (ppm) | 67.204 | 66.0667 | 68.3414 | 65.8228 | 68.5852
Ta, Tantalum (ppm) | 0.34833 | 0.28007 | 0.41658 | 0.31189 | 0.38477
Tb, Terbium (ppm) | 0.51813 | 0.48404 | 0.55221 | 0.49625 | 0.54
Te, Tellurium (ppm) | 0.049133 | 0.041278 | 0.056989 | IND | IND
Th, Thorium (ppm) | 8.9941 | 8.7383 | 9.2499 | 8.7673 | 9.2209
Ti, Titanium (wt.%) | 0.1609 | 0.1458 | 0.17599 | 0.15478 | 0.16701
Tl, Thallium (ppm) | 0.52737 | 0.51329 | 0.54145 | 0.50579 | 0.54894
Tm, Thulium (ppm) | 0.20233 | 0.1839 | 0.22076 | IND | IND
U, Uranium (ppm) | 13.7747 | 13.354 | 14.1954 | 13.4044 | 14.1451
V, Vanadium (ppm) | 112.2119 | 110.4297 | 113.9942 | 109.0986 | 115.3252
W, Tungsten (ppm) | 1.2697 | 1.1934 | 1.3461 | IND | IND
Y, Yttrium (ppm) | 12.9139 | 12.505 | 13.3228 | 12.418 | 13.4098
Yb, Ytterbium (ppm) | 1.4162 | 1.2912 | 1.5413 | 1.3286 | 1.5039
Zn, Zinc (ppm) | 55.2939 | 54.0369 | 56.5509 | 53.2281 | 57.3597
Zr, Zirconium (ppm) | 62.518 | 61.1116 | 63.9244 | 61.1181 | 63.9179
Aqua Regia Digestion
Ag, Silver (ppm) | 0.17772 | 0.17252 | 0.18292 | 0.16055 | 0.19488
Al, Aluminium (wt.%) | 0.5945 | 0.57033 | 0.61866 | 0.57345 | 0.61554
As, Arsenic (ppm) | 64.8996 | 63.419 | 66.3803 | 63.0932 | 66.7061
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 225.9637 | 214.0083 | 237.9191 | 217.6864 | 234.241
Be, Beryllium (ppm) | 0.62889 | 0.61378 | 0.644 | 0.59875 | 0.65903
Bi, Bismuth (ppm) | 2.5797 | 2.5299 | 2.6295 | 2.49 | 2.6693
Ca, Calcium (wt.%) | 3.8542 | 3.8015 | 3.9069 | 3.7877 | 3.9207
Cd, Cadmium (ppm) | 0.19873 | 0.19349 | 0.20398 | 0.17667 | 0.2208
Ce, Cerium (ppm) | 34.7548 | 32.7705 | 36.7391 | 33.8007 | 35.7088
Co, Cobalt (wt.%) | 0.76093 | 0.74103 | 0.78084 | 0.74502 | 0.77685
Cr, Chromium (ppm) | 15.5232 | 14.792 | 16.2544 | 14.0346 | 17.0118
Cs, Caesium (ppm) | 0.56649 | 0.50883 | 0.62415 | 0.53929 | 0.59369
Cu, Copper (wt.%) | 1.5785 | 1.5569 | 1.6002 | 1.5457 | 1.6114
Dy, Dysprosium (ppm) | 1.343 | 1.1544 | 1.5315 | IND | IND
Er, Erbium (ppm) | 0.57173 | 0.52949 | 0.61398 | IND | IND
Eu, Europium (ppm) | 0.4545 | 0.3468 | 0.5622 | IND | IND
Fe, Iron (wt.%) | 1.6391 | 1.6186 | 1.6596 | 1.6105 | 1.6677
Ga, Gallium (ppm) | 1.7642 | 1.6304 | 1.8979 | 1.6692 | 1.8592
Gd, Gadolinium (ppm) | 2.1755 | 1.573 | 2.7781 | 2.0073 | 2.3438
Hf, Hafnium (ppm) | 0.26654 | 0.24847 | 0.2846 | 0.25149 | 0.28158
Hg, Mercury (ppm) | 0.018636 | 0.015364 | 0.021908 | IND | IND
Ho, Holmium (ppm) | 0.22235 | 0.20078 | 0.24392 | IND | IND
In, Indium (ppm) | 0.063143 | 0.061988 | 0.064299 | 0.059978 | 0.066308
K, Potassium (wt.%) | 0.17697 | 0.16856 | 0.18538 | 0.17129 | 0.18265
La, Lanthanum (ppm) | 14.2363 | 13.2746 | 15.198 | 13.7402 | 14.7324
Li, Lithium (ppm) | 18.3578 | 17.0293 | 19.6863 | 17.5686 | 19.1471
Lu, Lutetium (ppm) | 0.079333 | 0.064656 | 0.09401 | IND | IND
Mg, Magnesium (wt.%) | 2.9015 | 2.8583 | 2.9446 | 2.8353 | 2.9676
Mn, Manganese (wt.%) | 0.10631 | 0.10414 | 0.10847 | 0.10444 | 0.10817
Mo, Molybdenum (ppm) | 3.3198 | 3.2455 | 3.394 | 3.2079 | 3.4316
Nd, Neodymium (ppm) | 12.9641 | 10.7341 | 15.1941 | 12.1881 | 13.74
Ni, Nickel (ppm) | 28.388 | 27.4386 | 29.3375 | 27.1433 | 29.6328
P, Phosphorus (wt.%) | 0.026448 | 0.025814 | 0.027083 | 0.025278 | 0.027619
Pb, Lead (ppm) | 10.9759 | 10.327 | 11.6248 | 10.509 | 11.4428
Pr, Praseodymium (ppm) | 2.8174 | 1.7119 | 3.9228 | 2.6362 | 2.9985
Rb, Rubidium (ppm) | 8.8308 | 8.321 | 9.3406 | 8.4298 | 9.2317
Re, Rhenium (ppm) | 0.044583 | 0.043389 | 0.045777 | 0.041621 | 0.047546
S, Sulphur (wt.%) | 0.74281 | 0.71816 | 0.76746 | 0.72821 | 0.75741
Sb, Antimony (ppm) | 1.4854 | 1.3848 | 1.586 | 1.4159 | 1.555
Sc, Scandium (ppm) | 2.2326 | 2.1534 | 2.3118 | 2.1151 | 2.3501
Se, Selenium (ppm) | 4.8965 | 4.703 | 5.09 | 4.3705 | 5.4224
Sn, Tin (ppm) | 0.31905 | 0.28877 | 0.34932 | 0.29049 | 0.3476
Sr, Strontium (ppm) | 40.2326 | 39.1909 | 41.2744 | 39.1839 | 41.2814
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.29164 | 0.26726 | 0.31601 | 0.26621 | 0.31707
Te, Tellurium (ppm) | 0.048241 | 0.044505 | 0.051978 | IND | IND
Th, Thorium (ppm) | 5.5111 | 5.3934 | 5.6287 | 5.3121 | 5.7101
Tl, Thallium (ppm) | 0.10886 | 0.10233 | 0.1154 | IND | IND
U, Uranium (ppm) | 11.8427 | 11.4704 | 12.215 | 11.4774 | 12.208
V, Vanadium (ppm) | 19.2956 | 18.5618 | 20.0294 | 18.3165 | 20.2747
W, Tungsten (ppm) | 0.10471 | 0.098412 | 0.111 | IND | IND
Y, Yttrium (ppm) | 5.6926 | 5.5092 | 5.876 | 5.5353 | 5.85
Yb, Ytterbium (ppm) | 0.49776 | 0.49357 | 0.50194 | IND | IND
Zn, Zinc (ppm) | 48.2215 | 46.7268 | 49.7161 | 46.2759 | 50.167
Zr, Zirconium (ppm) | 8.9917 | 8.7802 | 9.2031 | 8.6609 | 9.3224

## Indicative Values
Table 3. Indicative Values for OREAS 554
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
As | ppm | < 10 | Na2O | wt.% | 0.072151 | V2O5 | ppm | 207.6444
Bi | ppm | < 10 | Ni | ppm | 94.1347 | Zn | ppm | 66.1111
Cd | ppm | < 10 | Pb | ppm | 85.8333 | Zr | ppm | 66.6667
Cl | ppm | 255 | Sb | ppm | < 10
Mo | ppm | 28.6667 | Sn | ppm | < 10
Thermogravimetry
H2O- | wt.% | 0.12833
Sulphuric Acid 5% Leach
Cu | wt.% | 0.25038
Sulphuric Acid 10% Leach
Cu | wt.% | 0.51706
Peroxide Fusion ICP
Be | ppm | 2.2185 | Mo | ppm | 4.1837 | Te | ppm | < 1
Cd | ppm | 0.22778 | Nb | ppm | 7.8081 | Tl | ppm | 0.48429
Ge | ppm | 1.4542 | Re | ppm | < 0.1 | W | ppm | 1.5512
Hg | ppm | < 0.1 | Se | ppm | 13.343
In | ppm | < 0.2 | Sn | ppm | 2.4083
4-Acid Digestion
B | ppm | 7.3116 | Hg | ppm | 0.053333
Aqua Regia Digestion
Au | ppm | 0.01307 | Pd | ppb | < 10 | Ti | wt.% | 0.0032556
Ge | ppm | 0.045334 | Pt | ppb | 13.3333 | Tm | ppm | 0.085667
Na | wt.% | 0.017971 | Si | wt.% | 0.061083
Nb | ppm | < 0.05 | Sm | ppm | 2.5115

## Abbreviations
Table 4. Abbreviations used for OREAS 554
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with AAS finish
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digest with either AAS or ICP-OES finish as appropriate
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
AL-10S*AAS | acid leach of 10% sulphuric acid with AAS finish
AL-10S*OES | acid leach of 10% sulphuric acid with Optical Emmision Spectrometry [aka: A(tomic)ES, ICP-OES] finish
AL-5S*AAS | acid leach of 5% sulphuric acid with AAS finish
AL-5S*OES | acid leach of 5% sulphuric acid with OES finish
AL-5S*OES/AAS | acid leach of 5% sulphuric acid with either AAS or ICP-OES finish as appropriate
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/AAS | ore grade aqua regia digest with either AAS or ICP-OES finish as appropriate
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
IRC | infrared combustion furnace
LOI*Furnace | loss on ignition with muffle furnace finish
LOI*TGA | loss on ignition with Thermal Gravimetric Analyser finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
PSF*XRF | pyro-sulfate fusion with x-ray fluorescence finish
Text Values:
N.A. | Not Applicable (Lab 06)
N.A. | Not Applicable (Lab 21)
NA | Unable to report due to QC failure (Lab 18)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 554
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
AH Knight, Kitwe, Copperbelt, Zambia
AH Knight, Lumwana, Mwinilunga, Zambia
ALS, Brisbane, QLD, Australia
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Carsurin, Kendari, Sulawesi, Indonesia
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSALABS, Vancouver, BC, Canada
Nagrom, Perth, WA, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Robinson International Afrique SAS, Lubumbashi, Katanga, Congo, Democratic Republic of the (Zaire)
Ruashi Mine Site Laboratory, Lubumbashi, Katanga, Congo, Democratic Republic of the (Zaire)
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
UIS Analytical Services, Centurion , South Africa

## Fusion XRF
| Analytical results for Al2O3 in OREAS 554 (Certified Value 8.23 wt.%) | [20 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 8.36 | 8.48 | 8.41 | 8.24 | 8.21 | 8.106 | 8.18 | 8.3 | 8.13 | 8.29 | 8.19 | 8.44 | 8.2549 | 8.1 | 8.56 | 8.2333 | 8.24 | 8.1981 | 8.1 | 1
| 1 | 2 | 8.32 | 8.48 | 8.4 | 8.22 | 8.23 | 7.917 | 8.24 | 8.22 | 8.19 | 8.31 | 8.24 | 8.31 | 8.2051 | 8.23 | 7.94 | 8.1611 | 8.25 | 8.2661 | 8.09 | 0
| 1 | 3 | 8.3 | 8.5 | 8.37 | 8.22 | 8.23 | 7.86 | 8.18 | 8.28 | 8.17 | 8.23 | 8.16 | 8.33 | 8.1992 | 8.21 | 8.12 | 8.1552 | 8.28 | 8.1914 | 8.17 | 16
| 1 | 4 | 8.24 | 8.49 | 8.41 | 8.12 | 8.29 | 7.86 | 8.14 | 8.26 | 8.18 | 8.22 | 8.25 | 8.42 | 8.192 | 8.14 | 8.63 | 8.1747 | 8.34 | 8.1894 | 8.1 | 8.2301
| 1 | 5 | 8.3 | 8.49 | 8.43 | 8.15 | 8.23 | 7.936 | 8.19 | 8.27 | 8.21 | 8.23 | 8.31 | 8.36 | 8.3041 | 8.19 | 8.33 | 8.1585 | 8.31 | 8.263 | 8.14 | 7
| 1 | 6 | 8.23 | 8.46 | 8.37 | 8.06 | 8.31 | 7.993 | 8.15 | 8.22 | 8.11 | 8.2 | 8.13 | 8.45 | 8.2539 | 8.09 | 8.18 | 8.2095 | 8.29 | 8.2331 | 8.16
| Mean |  | 8.2917 | 8.4833 | 8.3983 | 8.1683 | 8.25 | 7.9453 | 8.18 | 8.2583 | 8.165 | 8.2467 | 8.2133 | 8.385 | 8.2348 | 8.16 | 8.2933 | 8.182 | 8.285 | 8.2235 | 8.1267
| Median |  | 8.3 | 8.485 | 8.405 | 8.185 | 8.23 | 7.9265 | 8.18 | 8.265 | 8.175 | 8.23 | 8.215 | 8.39 | 8.2295 | 8.165 | 8.255 | 8.1679 | 8.285 | 8.2156 | 8.12
| Std Dev. |  | 0.04916 | 0.013663 | 0.024014 | 0.070545 | 0.04 | 0.093316 | 0.035214 | 0.032506 | 0.037815 | 0.043205 | 0.065929 | 0.059582 | 0.043709 | 0.058652 | 0.26591 | 0.032097 | 0.037283 | 0.035517 | 0.034448
| Rel.Std.Dev. |  | 0.5929% | 0.1611% | 0.2859% | 0.8636% | 0.4848% | 1.174% | 0.4305% | 0.3936% | 0.4631% | 0.5239% | 0.8027% | 0.7106% | 0.5308% | 0.7188% | 3.206% | 0.3923% | 0.45% | 0.4319% | 0.4239%
| PDM3 |  | 0.7484% | 3.077% | 2.044% | -0.7502% | 0.2421% | -3.46% | -0.6084% | 0.3434% | -0.7907% | 0.2016% | -0.2034% | 1.882% | 0.05801% | -0.8514% | 0.7686% | -0.5836% | 0.6674% | -0.07947% | -1.256%
| Z-Score (Absolute) |  | 0.67 | 2.95 | 1.94 | 0.79 | 0.18 | 3.44 | 0.65 | 0.28 | 0.83 | 0.14 | 0.26 | 1.78 | 0 | 0.89 | 0.69 | 0.63 | 0.6 | 0.13 | 1.28
| Analytical results for As in OREAS 554 (Indicative Value < 10 ppm) | [20 empty] | Indicative
As | Round | Replicate | 02: Lab; 04: Lab | 1
| No | No | 02: 02; 04: 12 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF | 1
1
| 1 | 1 | 02: < 10; 04: < 10 | 1
| 1 | 2 | 02: < 10; 04: < 10 | 0
| 1 | 3 | 02: < 10; 04: < 10 | 16
| 1 | 4 | 02: < 10; 04: < 10 | < 10
| 1 | 5 | 02: < 10; 04: < 10 | 7
| 1 | 6 | 02: < 10; 04: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02 to 04
| Analytical results for BaO in OREAS 554 (Certified Value 1418 ppm) | [20 empty] | Certified
BaO | Round | Replicate | 02: Lab; 04: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab; 17: Lab | 1
| No | No | 02: 02; 04: 12; 09: 13; 10: 14; 11: 15; 12: 16; 13: 17; 14: 19; 15: 20; 16: 22; 17: 28 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF; 12: BF*XRF; 13: BF*XRF; 14: BF*XRF; 15: BF*XRF; 16: BF*XRF; 17: BF*XRF | 0
0
| 1 | 1 | 02: 1460; 04: 2713; 09: 1400; 10: 1500; 11: 1400; 12: 1500; 13: 1300; 14: 1476.9; 15: 1500; 16: 1467.4758; 17: 1300 | 1
| 1 | 2 | 02: 1430; 04: 2412; 09: 1400; 10: 1400; 11: 1500; 12: 1500; 13: 1300; 14: 1356.8; 15: 1400; 16: 1500.4812; 17: 1400 | 0
| 1 | 3 | 02: 1430; 04: 2423; 09: 1400; 10: 1500; 11: 1400; 12: 1400; 13: 1300; 14: 1445.8; 15: 1400; 16: 1507.4838; 17: 1400 | 16
| 1 | 4 | 02: 1450; 04: 2412; 09: 1400; 10: 1400; 11: 1400; 12: 1500; 13: 1300; 14: 1364; 15: 1500; 16: 1431.9009; 17: 1400 | 1417.9097
| 1 | 5 | 02: 1420; 04: 2434; 09: 1400; 10: 1400; 11: 1500; 12: 1500; 13: 1300; 14: 1303.2; 15: 1500; 16: 1461.753; 17: 1400 | 8
| 1 | 6 | 02: 1450; 04: 2479; 09: 1300; 10: 1400; 11: 1400; 12: 1500; 13: 1300; 14: 1278.4; 15: 1500; 16: 1440.3896; 17: 1500
| Mean |  | 02: 1440; 04: 2478.8333; 09: 1383.3333; 10: 1433.3333; 11: 1433.3333; 12: 1483.3333; 13: 1300; 14: 1370.85; 15: 1466.6667; 16: 1468.2474; 17: 1400
| Median |  | 02: 1440; 04: 2428.5; 09: 1400; 10: 1400; 11: 1400; 12: 1500; 13: 1300; 14: 1360.4; 15: 1500; 16: 1464.6144; 17: 1400
| Std Dev. |  | 02: 15.4919; 04: 117.3838; 09: 40.8248; 10: 51.6398; 11: 51.6398; 12: 40.8248; 13: 0; 14: 77.7585; 15: 51.6398; 16: 30.7194; 17: 63.2456
| Rel.Std.Dev. |  | 02: 1.076%; 04: 4.735%; 09: 2.951%; 10: 3.603%; 11: 3.603%; 12: 2.752%; 13: 0%; 14: 5.672%; 15: 3.521%; 16: 2.092%; 17: 4.518%
| PDM3 |  | 02: 1.558%; 04: 74.82%; 09: -2.439%; 10: 1.088%; 11: 1.088%; 12: 4.614%; 13: -8.316%; 14: -3.319%; 15: 3.439%; 16: 3.55%; 17: -1.263%
| Z-Score (Absolute) |  | 02: 0.13; 04: 20.19; 09: 0.97; 10: 0; 11: 0; 12: 0.97; 13: 2.58; 14: 1.21; 15: 0.64; 16: 0.67; 17: 0.64
| Analytical results for Bi in OREAS 554 (Indicative Value < 10 ppm) | [20 empty] | Indicative
Bi | Round | Replicate | 02: Lab | 1
| No | No | 02: 12 | ppm
|  |  | 02: BF*XRF | 1
1
| 1 | 1 | 02: < 10 | 1
| 1 | 2 | 02: < 10 | 0
| 1 | 3 | 02: < 10 | 16
| 1 | 4 | 02: < 10 | < 10
| 1 | 5 | 02: < 10 | 8
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for CaO in OREAS 554 (Certified Value 5.58 wt.%) | [20 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 5.57 | 5.66 | 5.6 | 5.65 | 5.59 | 5.261 | 5.65 | 5.52 | 5.56 | 5.53 | 5.56 | 5.47 | 5.7293 | 5.48 | 5.64 | 5.5798 | 5.64 | 5.2631 | 5.53 | 1
| 1 | 2 | 5.61 | 5.66 | 5.6 | 5.58 | 5.59 | 5.275 | 5.64 | 5.46 | 5.6 | 5.56 | 5.59 | 5.36 | 5.7242 | 5.53 | 5.36 | 5.5593 | 5.6 | 5.282 | 5.52 | 0
| 1 | 3 | 5.58 | 5.65 | 5.61 | 5.61 | 5.62 | 5.261 | 5.64 | 5.52 | 5.6 | 5.55 | 5.55 | 5.42 | 5.6546 | 5.49 | 5.52 | 5.5526 | 5.64 | 5.263 | 5.51 | 16
| 1 | 4 | 5.57 | 5.65 | 5.63 | 5.59 | 5.59 | 5.317 | 5.61 | 5.5 | 5.6 | 5.54 | 5.62 | 5.45 | 5.7139 | 5.42 | 5.68 | 5.5669 | 5.61 | 5.2863 | 5.5 | 5.5801
| 1 | 5 | 5.58 | 5.65 | 5.63 | 5.63 | 5.64 | 5.233 | 5.62 | 5.52 | 5.64 | 5.53 | 5.52 | 5.43 | 5.7138 | 5.47 | 5.71 | 5.5503 | 5.58 | 5.2693 | 5.5 | 9
| 1 | 6 | 5.58 | 5.65 | 5.61 | 5.69 | 5.6 | 5.317 | 5.63 | 5.46 | 5.57 | 5.57 | 5.52 | 5.46 | 5.7527 | 5.48 | 5.63 | 5.5728 | 5.61 | 5.2863 | 5.53
| Mean |  | 5.5817 | 5.6533 | 5.6133 | 5.625 | 5.605 | 5.2773 | 5.6317 | 5.4967 | 5.595 | 5.5467 | 5.56 | 5.4317 | 5.7148 | 5.4783 | 5.59 | 5.5636 | 5.6133 | 5.275 | 5.515
| Median |  | 5.58 | 5.65 | 5.61 | 5.62 | 5.595 | 5.268 | 5.635 | 5.51 | 5.6 | 5.545 | 5.555 | 5.44 | 5.7191 | 5.48 | 5.635 | 5.5631 | 5.61 | 5.2756 | 5.515
| Std Dev. |  | 0.01472 | 0.005164 | 0.013663 | 0.040866 | 0.020736 | 0.033619 | 0.01472 | 0.029439 | 0.028107 | 0.01633 | 0.039497 | 0.039707 | 0.032731 | 0.035449 | 0.12992 | 0.011634 | 0.023381 | 0.011158 | 0.013784
| Rel.Std.Dev. |  | 0.2637% | 0.09134% | 0.2434% | 0.7265% | 0.37% | 0.6371% | 0.2614% | 0.5356% | 0.5024% | 0.2944% | 0.7104% | 0.731% | 0.5727% | 0.6471% | 2.324% | 0.2091% | 0.4165% | 0.2115% | 0.2499%
| PDM3 |  | 0.02876% | 1.313% | 0.5963% | 0.8053% | 0.4469% | -5.425% | 0.9248% | -1.495% | 0.2677% | -0.5985% | -0.3595% | -2.659% | 2.414% | -1.823% | 0.1781% | -0.2946% | 0.5963% | -5.467% | -1.166%
| Z-Score (Absolute) |  | 0 | 1.12 | 0.49 | 0.67 | 0.36 | 4.74 | 0.78 | 1.32 | 0.21 | 0.54 | 0.34 | 2.33 | 2.07 | 1.61 | 0.13 | 0.28 | 0.49 | 4.77 | 1.04
| Analytical results for Cd in OREAS 554 (Indicative Value < 10 ppm) | [20 empty] | Indicative
Cd | Round | Replicate | 02: Lab | 1
| No | No | 02: 12 | ppm
|  |  | 02: BF*XRF | 1
1
| 1 | 1 | 02: < 10 | 1
| 1 | 2 | 02: < 10 | 0
| 1 | 3 | 02: < 10 | 16
| 1 | 4 | 02: < 10 | < 10
| 1 | 5 | 02: < 10 | 9
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for Cl in OREAS 554 (Indicative Value 255 ppm) | [20 empty] | Indicative
Cl | Round | Replicate | 02: Lab; 04: Lab; 09: Lab | 1
| No | No | 02: 02; 04: 12; 09: 23 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF | 0
0
| 1 | 1 | 02: < 20; 04: 320; 09: 190 | 1
| 1 | 2 | 02: < 20; 04: 310; 09: 170 | 0
| 1 | 3 | 02: < 20; 04: 310; 09: 180 | 16
| 1 | 4 | 02: < 20; 04: 320; 09: 240 | 255
| 1 | 5 | 02: < 20; 04: 310; 09: 210 | 10
| 1 | 6 | 02: < 20; 04: 310; 09: 190
| Mean |  | 04: 313.3333; 09: 196.6667
| Median |  | 04: 310; 09: 190
| Std Dev. |  | 04: 5.164; 09: 25.0333
| Rel.Std.Dev. |  | 04: 1.648%; 09: 12.73%
| PDM3 |  | 04: 22.88%; 09: -22.88%
| Z-Score (Absolute) |  | 02: 1.08; 04: 0.67; 09: 0
| Analytical results for Co in OREAS 554 (Certified Value 0.805 wt.%) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 19 | 21 | 23 | 25 | [6 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.815 | 0.799 | 0.814 | 0.795 | 0.8 | 0.659 | 0.834 | 0.837 | 0.802 | 0.78946 | 0.85 | 0.795 | 0.8153 | [6 empty] | 1
| 1 | 2 | 0.807 | 0.794 | 0.815 | 0.773 | 0.79 | 0.547 | 0.834 | 0.845 | 0.802 | 0.7807 | 0.78 | 0.766 | 0.81416 | [6 empty] | 0
| 1 | 3 | 0.807 | 0.805 | 0.818 | 0.771 | 0.79 | 0.527 | 0.81 | 0.832 | 0.794 | 0.78666 | 0.79 | 0.767 | 0.81336 | [6 empty] | 16
| 1 | 4 | 0.82 | 0.807 | 0.83 | 0.76 | 0.8 | 0.535 | 0.818 | 0.858 | 0.802 | 0.7924 | 0.81 | 0.78 | 0.81616 | [6 empty] | 0.80452
| 1 | 5 | 0.808 | 0.808 | 0.818 | 0.763 | 0.8 | 0.535 | 0.818 | 0.842 | 0.818 | 0.80068 | 0.81 | 0.769 | 0.81958 | [6 empty] | 10
| 1 | 6 | 0.816 | 0.801 | 0.823 | 0.777 | 0.8 | 0.514 | 0.818 | 0.818 | 0.802 | 0.79983 | 0.79 | 0.783 | 0.81673
| Mean |  | 0.81217 | 0.80233 | 0.81967 | 0.77317 | 0.79667 | 0.55283 | 0.822 | 0.83867 | 0.80333 | 0.79162 | 0.805 | 0.77667 | 0.81588
| Median |  | 0.8115 | 0.803 | 0.818 | 0.772 | 0.8 | 0.535 | 0.818 | 0.8395 | 0.802 | 0.79093 | 0.8 | 0.7745 | 0.81573
| Std Dev. |  | 0.0055648 | 0.0053541 | 0.0059554 | 0.012432 | 0.005164 | 0.053135 | 0.009798 | 0.013411 | 0.0078655 | 0.0077291 | 0.0251 | 0.011431 | 0.0021961
| Rel.Std.Dev. |  | 0.6852% | 0.6673% | 0.7266% | 1.608% | 0.6482% | 9.611% | 1.192% | 1.599% | 0.9791% | 0.9764% | 3.118% | 1.472% | 0.2692%
| PDM3 |  | 0.9505% | -0.2717% | 1.883% | -3.897% | -0.9761% | -31.28% | 2.173% | 4.244% | -0.1474% | -1.603% | 0.05974% | -3.462% | 1.412%
| Z-Score (Absolute) |  | 0.47 | 0.05 | 0.88 | 1.62 | 0.36 | 13.46 | 1 | 1.9 | 0 | 0.63 | 0.09 | 1.43 | 0.67
| Analytical results for Cr2O3 in OREAS 554 (Certified Value 97 ppm) | [20 empty] | Certified
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 90 | 90 | 300 | < 100 | 58 | < 100 | 100 | < 100 | 100 | 100 | < 100 | 151.6 | < 100 | < 100 | 115.8534 | 120 | < 100 | < 100 |  | 1
| 1 | 2 | 80 | 90 | 100 | < 100 | 44 | 100 | 100 | < 100 | 100 | 100 | < 100 | 162.4 | < 100 | < 100 | 121.1513 | < 50 | < 100 | < 100 |  | 0
| 1 | 3 | 100 | 90 | 200 | < 100 | 102 | < 100 | 100 | 400 | 100 | 200 | < 100 | 136.8 | 100 | < 100 | 94.1617 | 130 | < 100 | < 100 |  | 16
| 1 | 4 | 80 | 100 | 100 | < 100 | 58 | < 100 | 100 | 300 | 100 | < 100 | < 100 | 190 | < 100 | 100 | 114.8392 | < 50 | < 100 | < 100 |  | 97.1022
| 1 | 5 | 90 | 90 | < 100 | < 100 | 58 | 100 | 100 | 100 | 100 | < 100 | < 100 | 242.7 | < 100 | < 100 | 130.93 | 110 | < 100 | < 100 |  | 11
| 1 | 6 | 90 | 80 | 100 | < 100 | 58 | < 100 | 200 | < 100 | 100 | 100 | < 100 | 165.1 | 100 | < 100 | 108.5901 | < 50 | < 100 | < 100
| Mean |  | 88.3333 | 90 | 160 |  | 63 | 100 | 116.6667 | 266.6667 | 100 | 125 |  | 174.7667 | 100 | 100 | 114.2543 | 120
| Median |  | 90 | 90 | 100 |  | 58 | 100 | 100 | 300 | 100 | 100 |  | 163.75 | 100 | 100 | 115.3463 | 120
| Std Dev. |  | 7.5277 | 6.3246 | 89.4427 |  | 19.9098 | 0 | 40.8248 | 152.7525 | 0 | 50 |  | 37.6073 | 0 |  | 12.364 | 10
| Rel.Std.Dev. |  | 8.522% | 7.027% | 55.9% |  | 31.6% | 0% | 34.99% | 57.28% | 0% | 40% |  | 21.52% | 0% |  | 10.82% | 8.333%
| PDM3 |  | -9.031% | -7.314% | 64.77% |  | -35.12% | 2.984% | 20.15% | 174.6% | 2.984% | 28.73% |  | 79.98% | 2.984% | 2.984% | 17.66% | 23.58%
| Z-Score (Absolute) |  | 0.2 | 0.25 | 1.57 | 0.78 | 0.44 | 0.35 | 0.93 | 2 | 0.5 | 0.5 | 0.78 | 2.42 | 0.35 | 0.57 | 0.87 | 0.2 | 0.78 | 0.78
| Analytical results for Cu in OREAS 554 (Certified Value 1.58 wt.%) | [20 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 18 | 19 | 21 | 23 | 25 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | PSF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 2
3
| 1 | 1 | 1.601 | 1.635 | 1.594 | 1.5578 | 1.58 | 1.738 | > 1.6 | > 1.6 | 1.565 | 1.59 | 1.5635 | 1.59 | 1.511 | 1.5527 | [5 empty] | 1
| 1 | 2 | 1.602 | 1.635 | 1.596 | 1.5418 | 1.59 | 1.767 | > 1.6 | > 1.6 | 1.58 | 1.534 | 1.5579 | 1.48 | 1.612 | 1.5632 | [5 empty] | 0
| 1 | 3 | 1.588 | 1.635 | 1.568 | 1.5338 | 1.61 | 1.727 | > 1.6 | > 1.6 | 1.575 | 1.558 | 1.5817 | 1.53 | 1.436 | 1.55 | [5 empty] | 16
| 1 | 4 | 1.585 | 1.635 | 1.567 | 1.5178 | 1.57 | 1.71 | > 1.6 | > 1.6 | 1.575 | 1.574 | 1.5972 | 1.58 | 1.417 | 1.5516 | [5 empty] | 1.5786
| 1 | 5 | 1.582 | 1.645 | 1.591 | 1.5338 | 1.56 | 1.69 | > 1.6 | > 1.6 | 1.565 | 1.574 | 1.5927 | 1.58 | 1.448 | 1.5612 | [5 empty] | 12
| 1 | 6 | 1.603 | 1.635 | 1.587 | 1.5578 | 1.6 | 1.72 | > 1.6 | > 1.6 | 1.58 | 1.574 | 1.5584 | 1.56 | 1.444 | 1.5568
| Mean |  | 1.5935 | 1.6367 | 1.5838 | 1.5405 | 1.585 | 1.7253 |  |  | 1.5733 | 1.5673 | 1.5752 | 1.5533 | 1.478 | 1.5559
| Median |  | 1.5945 | 1.635 | 1.589 | 1.5378 | 1.585 | 1.7235 |  |  | 1.575 | 1.574 | 1.5726 | 1.57 | 1.446 | 1.5548
| Std Dev. |  | 0.0095237 | 0.0040825 | 0.013014 | 0.015526 | 0.018708 | 0.02612 |  |  | 0.0068313 | 0.019211 | 0.017607 | 0.041793 | 0.072919 | 0.0053993
| Rel.Std.Dev. |  | 0.5977% | 0.2494% | 0.8217% | 1.008% | 1.18% | 1.514% |  |  | 0.4342% | 1.226% | 1.118% | 2.691% | 4.934% | 0.347%
| PDM3 |  | 0.9461% | 3.681% | 0.3337% | -2.413% | 0.4076% | 9.298% |  |  | -0.3314% | -0.7115% | -0.2107% | -1.598% | -6.371% | -1.436%
| Z-Score (Absolute) |  | 0.69 | 2.24 | 0.34 | 1.21 | 0.38 | 5.42 | NA | NA | 0.03 | 0.26 | 0.03 | 0.75 | 3.45 | 0.66
| Analytical results for Fe2O3 in OREAS 554 (Certified Value 2.57 wt.%) | [20 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.573 | 2.645 | 2.58 | 2.63 | 2.59 | 2.588 | 2.6 | 2.57 | 2.59 | 2.58 | 2.54 | 2.83 | 2.5558 | 2.55 | 2.559 | 2.7114 | 2.56 | 2.4993 | 2.57 | 1
| 1 | 2 | 2.573 | 2.688 | 2.58 | 2.58 | 2.6 | 2.588 | 2.62 | 2.55 | 2.64 | 2.57 | 2.55 | 2.78 | 2.5458 | 2.55 | 2.488 | 2.6724 | 2.56 | 2.4988 | 2.56 | 0
| 1 | 3 | 2.573 | 2.631 | 2.57 | 2.62 | 2.59 | 2.588 | 2.59 | 2.57 | 2.64 | 2.59 | 2.53 | 2.77 | 2.5394 | 2.57 | 2.488 | 2.671 | 2.56 | 2.4911 | 2.57 | 16
| 1 | 4 | 2.588 | 2.645 | 2.6 | 2.59 | 2.61 | 2.616 | 2.57 | 2.55 | 2.63 | 2.6 | 2.56 | 2.78 | 2.5435 | 2.54 | 2.559 | 2.6837 | 2.57 | 2.5157 | 2.55 | 2.5749
| 1 | 5 | 2.588 | 2.645 | 2.59 | 2.59 | 2.61 | 2.588 | 2.59 | 2.57 | 2.64 | 2.57 | 2.58 | 2.8 | 2.5674 | 2.55 | 2.573 | 2.6745 | 2.57 | 2.5109 | 2.57 | 13
| 1 | 6 | 2.602 | 2.645 | 2.56 | 2.6 | 2.59 | 2.588 | 2.58 | 2.54 | 2.59 | 2.59 | 2.52 | 2.81 | 2.5688 | 2.54 | 2.545 | 2.6902 | 2.56 | 2.5099 | 2.59
| Mean |  | 2.5828 | 2.6498 | 2.58 | 2.6017 | 2.5983 | 2.5927 | 2.5917 | 2.5583 | 2.6217 | 2.5833 | 2.5467 | 2.795 | 2.5534 | 2.55 | 2.5353 | 2.6839 | 2.5633 | 2.5043 | 2.5683
| Median |  | 2.5805 | 2.645 | 2.58 | 2.595 | 2.595 | 2.588 | 2.59 | 2.56 | 2.635 | 2.585 | 2.545 | 2.79 | 2.5508 | 2.55 | 2.552 | 2.6791 | 2.56 | 2.5046 | 2.57
| Std Dev. |  | 0.011923 | 0.019518 | 0.014142 | 0.019408 | 0.0098319 | 0.011431 | 0.017224 | 0.013292 | 0.024833 | 0.012111 | 0.021602 | 0.022583 | 0.01259 | 0.010954 | 0.037718 | 0.015392 | 0.005164 | 0.0093323 | 0.013292
| Rel.Std.Dev. |  | 0.4616% | 0.7366% | 0.5481% | 0.746% | 0.3784% | 0.4409% | 0.6646% | 0.5195% | 0.9472% | 0.4688% | 0.8483% | 0.808% | 0.4931% | 0.4296% | 1.488% | 0.5735% | 0.2015% | 0.3727% | 0.5175%
| PDM3 |  | 0.3066% | 2.909% | 0.1965% | 1.038% | 0.9085% | 0.6884% | 0.6496% | -0.6449% | 1.815% | 0.326% | -1.098% | 8.546% | -0.8348% | -0.9686% | -1.538% | 4.23% | -0.4507% | -2.745% | -0.2566%
| Z-Score (Absolute) |  | 0 | 1.82 | 0.08 | 0.51 | 0.42 | 0.26 | 0.24 | 0.67 | 1.06 | 0.01 | 0.99 | 5.8 | 0.81 | 0.9 | 1.3 | 2.76 | 0.54 | 2.15 | 0.4
| Analytical results for K2O in OREAS 554 (Certified Value 2.01 wt.%) | [20 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.03 | 2.05 | 2.028 | 2.02 | 2.01 | 2.0213 | 2.03 | 2.01 | 1.97 | 2.02 | 2.02 | 2 | 1.9459 | 2 | 2.07 | 2.0528 | 1.98 | 1.9467 | 2.02 | 1
| 1 | 2 | 2.02 | 2.05 | 2.031 | 1.97 | 2.01 | 2.0177 | 2.06 | 1.99 | 1.99 | 2.01 | 2.02 | 1.98 | 1.9548 | 1.99 | 1.99 | 2.0388 | 1.98 | 1.9539 | 2.01 | 0
| 1 | 3 | 2.02 | 2.05 | 2.017 | 2.01 | 2.01 | 2.0261 | 2.05 | 2.01 | 1.97 | 2.01 | 2 | 1.98 | 1.9672 | 2 | 1.98 | 2.0413 | 1.99 | 1.9511 | 2.01 | 16
| 1 | 4 | 2.02 | 2.05 | 2.035 | 1.97 | 2.02 | 2.0273 | 2.03 | 2 | 1.98 | 2.01 | 2.03 | 1.99 | 1.9765 | 1.98 | 2.01 | 2.026 | 2 | 1.9574 | 2.01 | 2.0056
| 1 | 5 | 2.02 | 2.04 | 2.035 | 2 | 2.02 | 2.049 | 2.03 | 2.01 | 2.02 | 2.01 | 2.03 | 2 | 1.9618 | 1.99 | 2 | 2.0368 | 1.99 | 1.9487 | 2.01 | 14
| 1 | 6 | 2.01 | 2.05 | 2.025 | 2 | 2.01 | 2.0057 | 2.02 | 1.99 | 1.98 | 2.02 | 2 | 1.99 | 1.962 | 1.99 | 1.99 | 2.0453 | 1.99 | 1.9538 | 2.02
| Mean |  | 2.02 | 2.0483 | 2.0285 | 1.995 | 2.0133 | 2.0245 | 2.0367 | 2.0017 | 1.985 | 2.0133 | 2.0167 | 1.99 | 1.9614 | 1.9917 | 2.0067 | 2.0401 | 1.9883 | 1.952 | 2.0133
| Median |  | 2.02 | 2.05 | 2.0295 | 2 | 2.01 | 2.0237 | 2.03 | 2.005 | 1.98 | 2.01 | 2.02 | 1.99 | 1.9619 | 1.99 | 1.995 | 2.04 | 1.99 | 1.9525 | 2.01
| Std Dev. |  | 0.0063246 | 0.0040825 | 0.0068629 | 0.020736 | 0.005164 | 0.014287 | 0.015055 | 0.0098319 | 0.018708 | 0.005164 | 0.013663 | 0.0089443 | 0.010458 | 0.0075277 | 0.03266 | 0.0089417 | 0.0075277 | 0.0038991 | 0.005164
| Rel.Std.Dev. |  | 0.3131% | 0.1993% | 0.3383% | 1.039% | 0.2565% | 0.7057% | 0.7392% | 0.4912% | 0.9425% | 0.2565% | 0.6775% | 0.4495% | 0.5332% | 0.378% | 1.628% | 0.4383% | 0.3786% | 0.1998% | 0.2565%
| PDM3 |  | 0.7168% | 2.13% | 1.141% | -0.5297% | 0.3844% | 0.942% | 1.548% | -0.1973% | -1.028% | 0.3844% | 0.5506% | -0.779% | -2.207% | -0.6959% | 0.05205% | 1.721% | -0.8621% | -2.676% | 0.3844%
| Z-Score (Absolute) |  | 0.25 | 1.29 | 0.56 | 0.67 | 0 | 0.41 | 0.86 | 0.43 | 1.04 | 0 | 0.12 | 0.86 | 1.91 | 0.8 | 0.25 | 0.99 | 0.92 | 2.26 | 0
| Analytical results for MgO in OREAS 554 (Certified Value 5.27 wt.%) | [20 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 5.24 | 5.23 | 5.36 | 5.38 | 5.32 | 4.908 | 5.27 | 5.3 | 5.27 | 5.23 | 5.27 | NA | 5.4764 | 5.23 | 5.37 | 5.2988 | 5.35 | 4.5925 | 5.15 | 1
| 1 | 2 | 5.25 | 5.23 | 5.37 | 5.31 | 5.34 | 4.941 | 5.27 | 5.27 | 5.29 | 5.26 | 5.3 | NA | 5.5183 | 5.34 | 5.08 | 5.2827 | 5.32 | 4.5844 | 5.13 | 0
| 1 | 3 | 5.23 | 5.25 | 5.38 | 5.31 | 5.31 | 4.908 | 5.26 | 5.31 | 5.28 | 5.25 | 5.26 | NA | 5.4505 | 5.24 | 4.97 | 5.2672 | 5.33 | 4.6037 | 5.16 | 16
| 1 | 4 | 5.23 | 5.24 | 5.4 | 5.31 | 5.31 | 4.908 | 5.24 | 5.29 | 5.31 | 5.26 | 5.31 | NA | 5.5343 | 5.26 | 4.96 | 5.2684 | 5.28 | 4.597 | 5.17 | 5.2726
| 1 | 5 | 5.24 | 5.23 | 5.39 | 5.3 | 5.34 | 4.925 | 5.25 | 5.29 | 5.33 | 5.24 | 5.25 | NA | 5.5459 | 5.27 | 5.24 | 5.2502 | 5.29 | 4.6006 | 5.13 | 15
| 1 | 6 | 5.25 | 5.23 | 5.37 | 5.39 | 5.34 | 4.925 | 5.27 | 5.25 | 5.25 | 5.24 | 5.23 | NA | 5.4846 | 5.18 | 5.29 | 5.2667 | 5.28 | 4.6051 | 5.14
| Mean |  | 5.24 | 5.235 | 5.3783 | 5.3333 | 5.3267 | 4.9192 | 5.26 | 5.285 | 5.2883 | 5.2467 | 5.27 |  | 5.5017 | 5.2533 | 5.1517 | 5.2724 | 5.3083 | 4.5972 | 5.1467
| Median |  | 5.24 | 5.23 | 5.375 | 5.31 | 5.33 | 4.9165 | 5.265 | 5.29 | 5.285 | 5.245 | 5.265 |  | 5.5014 | 5.25 | 5.16 | 5.2678 | 5.305 | 4.5988 | 5.145
| Std Dev. |  | 0.0089443 | 0.0083666 | 0.01472 | 0.040332 | 0.015055 | 0.013556 | 0.012649 | 0.021679 | 0.028577 | 0.012111 | 0.030332 |  | 0.036991 | 0.052789 | 0.17291 | 0.016564 | 0.029269 | 0.0077722 | 0.01633
| Rel.Std.Dev. |  | 0.1707% | 0.1598% | 0.2737% | 0.7562% | 0.2826% | 0.2756% | 0.2405% | 0.4102% | 0.5404% | 0.2308% | 0.5756% |  | 0.6724% | 1.005% | 3.356% | 0.3142% | 0.5514% | 0.1691% | 0.3173%
| PDM3 |  | -0.6183% | -0.7131% | 2.005% | 1.152% | 1.025% | -6.703% | -0.239% | 0.2352% | 0.2984% | -0.4919% | -0.04933% |  | 4.344% | -0.3654% | -2.294% | -0.004731% | 0.6777% | -12.81% | -2.388%
| Z-Score (Absolute) |  | 0.46 | 0.55 | 2.08 | 1.26 | 1.13 | 6.36 | 0.09 | 0.37 | 0.43 | 0.34 | 0.09 | NA | 4.35 | 0.21 | 2.08 | 0.14 | 0.8 | 12.28 | 2.18
| Analytical results for MnO in OREAS 554 (Certified Value 0.144 wt.%) | [20 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.15 | 0.15 | 0.142 | 0.145 | 0.15 | 0.2001 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14404 | 0.14 | 0.14 | 0.15348 | 0.14 | 0.13398 | 0.14 | 1
| 1 | 2 | 0.14 | 0.15 | 0.142 | 0.144 | 0.15 | 0.1988 | 0.15 | 0.14 | 0.15 | 0.14 | 0.14 | 0.14 | 0.1447 | 0.14 | 0.13 | 0.15696 | 0.14 | 0.1331 | 0.14 | 0
| 1 | 3 | 0.14 | 0.15 | 0.142 | 0.147 | 0.16 | 0.1924 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.15 | 0.14668 | 0.14 | 0.14 | 0.15147 | 0.14 | 0.13302 | 0.15 | 16
| 1 | 4 | 0.15 | 0.15 | 0.144 | 0.144 | 0.15 | 0.1911 | 0.15 | 0.14 | 0.15 | 0.14 | 0.14 | 0.14 | 0.15038 | 0.14 | 0.15 | 0.15458 | 0.14 | 0.1332 | 0.15 | 0.14408
| 1 | 5 | 0.15 | 0.15 | 0.142 | 0.144 | 0.15 | 0.1924 | 0.15 | 0.14 | 0.15 | 0.14 | 0.14 | 0.15 | 0.15038 | 0.14 | 0.14 | 0.15407 | 0.14 | 0.1332 | 0.14 | 16
| 1 | 6 | 0.14 | 0.15 | 0.143 | 0.148 | 0.16 | 0.195 | 0.15 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14811 | 0.14 | 0.14 | 0.15885 | 0.14 | 0.13369 | 0.14
| Mean |  | 0.145 | 0.15 | 0.1425 | 0.14533 | 0.15333 | 0.19497 | 0.15 | 0.14 | 0.145 | 0.14 | 0.14 | 0.14333 | 0.14738 | 0.14 | 0.14 | 0.1549 | 0.14 | 0.13336 | 0.14333
| Median |  | 0.145 | 0.15 | 0.142 | 0.1445 | 0.15 | 0.1937 | 0.15 | 0.14 | 0.145 | 0.14 | 0.14 | 0.14 | 0.1474 | 0.14 | 0.14 | 0.15433 | 0.14 | 0.1332 | 0.14
| Std Dev. |  | 0.0054772 | 0 | 0.00083666 | 0.0017512 | 0.005164 | 0.0037195 | 0 | 0 | 0.0054772 | 0 | 0 | 0.005164 | 0.0027311 | 0 | 0.0063246 | 0.0026233 | 0 | 0.00038227 | 0.005164
| Rel.Std.Dev. |  | 3.777% | 0% | 0.5871% | 1.205% | 3.368% | 1.908% | 0% | 0% | 3.777% | 0% | 0% | 3.603% | 1.853% | 0% | 4.518% | 1.694% | 0% | 0.2866% | 3.603%
| PDM3 |  | 0.637% | 4.107% | -1.098% | 0.8683% | 6.421% | 35.32% | 4.107% | -2.833% | 0.637% | -2.833% | -2.833% | -0.5198% | 2.29% | -2.833% | -2.833% | 7.509% | -2.833% | -7.439% | -0.5198%
| Z-Score (Absolute) |  | 0.34 | 1.35 | 0.17 | 0.4 | 2.02 | 10.45 | 1.35 | 0.67 | 0.34 | 0.67 | 0.67 | 0 | 0.82 | 0.67 | 0.67 | 2.34 | 0.67 | 2.02 | 0
| Analytical results for Mo in OREAS 554 (Indicative Value 28.7 ppm) | [20 empty] | Indicative
Mo | Round | Replicate | 02: Lab | 1
| No | No | 02: 12 | ppm
|  |  | 02: BF*XRF | 1
1
| 1 | 1 | 02: 35 | 1
| 1 | 2 | 02: 26 | 0
| 1 | 3 | 02: 24 | 16
| 1 | 4 | 02: 30 | 28.6667
| 1 | 5 | 02: 27 | 11
| 1 | 6 | 02: 30
| Mean |  | 02: 28.6667
| Median |  | 02: 28.5
| Std Dev. |  | 02: 3.8816
| Rel.Std.Dev. |  | 02: 13.54%
| PDM3 |  | 02: -1.11e-13%
Z-Score (Absolute) | bar: 02
| Analytical results for Na2O in OREAS 554 (Indicative Value 0.072 wt.%) | [20 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 25 | 28 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.08 | 0.05 | < 0.05 | 0.05 | 0.13 | 0.05 | 0.1 | 0.08 | < 0.1 | 0.03 | 0.086935 | 0.12 | 0.012582 | 0.08 | [5 empty] | 1
| 1 | 2 | 0.08 | 0.05 | < 0.05 | 0.05 | 0.13 | 0.06 | 0.1 | 0.08 | < 0.1 | 0.02 | 0.10374 | 0.11 | 0.012681 | 0.08 | [5 empty] | 0
| 1 | 3 | 0.07 | 0.04 | < 0.05 | 0.06 | 0.13 | 0.05 | 0.1 | 0.07 | < 0.1 | 0.03 | 0.15111 | 0.1 | 0.012557 | 0.07 | [5 empty] | 16
| 1 | 4 | 0.08 | 0.05 | < 0.05 | 0.05 | 0.13 | 0.05 | 0.1 | 0.08 | < 0.1 | 0.03 | 0.10721 | 0.12 | 0.012582 | 0.07 | [5 empty] | 0.072151
| 1 | 5 | 0.08 | 0.06 | < 0.05 | 0.05 | 0.13 | 0.06 | 0.11 | 0.07 | < 0.1 | 0.04 | 0.080352 | 0.12 | 0.012582 | 0.09 | [5 empty] | 7
| 1 | 6 | 0.08 | 0.05 | < 0.05 | 0.05 | 0.13 | 0.05 | 0.1 | 0.06 | < 0.1 | 0.04 | 0.087859 | 0.1 | 0.012582 | 0.08
| Mean |  | 0.078333 | 0.05 |  | 0.051667 | 0.13 | 0.053333 | 0.10167 | 0.073333 |  | 0.031667 | 0.10287 | 0.11167 | 0.012595 | 0.078333
| Median |  | 0.08 | 0.05 |  | 0.05 | 0.13 | 0.05 | 0.1 | 0.075 |  | 0.03 | 0.0958 | 0.115 | 0.012582 | 0.08
| Std Dev. |  | 0.0040825 | 0.0063246 |  | 0.0040825 | 0 | 0.005164 | 0.0040825 | 0.008165 |  | 0.0075277 | 0.025818 | 0.0098319 | 4.3445e-05 | 0.0075277
| Rel.Std.Dev. |  | 5.212% | 12.65% |  | 7.902% | 0% | 9.682% | 4.016% | 11.13% |  | 23.77% | 25.1% | 8.805% | 0.345% | 9.61%
| PDM3 |  | 8.568% | -30.7% |  | -28.39% | 80.18% | -26.08% | 40.91% | 1.638% |  | -56.11% | 42.57% | 54.77% | -82.54% | 8.568%
| Z-Score (Absolute) |  | 0.43 | 0.39 | 1.11 | 0.34 | 1.93 | 0.29 | 1.11 | 0.29 | 0.39 | 0.92 | 1.14 | 1.4 | 1.47 | 0.43
| Analytical results for Ni in OREAS 554 (Indicative Value 94 ppm) | [20 empty] | Indicative
Ni | Round | Replicate | 02: Lab; 04: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab | 1
| No | No | 02: 02; 04: 09; 09: 10; 10: 19; 11: 21; 12: 23; 13: 25 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF; 12: BF*XRF; 13: BF*XRF | 0
0
| 1 | 1 | 02: 30; 04: 50; 09: < 24; 10: 100.28; 11: N.A.; 12: 200; 13: < 100 | 1
| 1 | 2 | 02: 40; 04: 50; 09: < 24; 10: 85.5; 11: N.A.; 12: 190; 13: < 100 | 0
| 1 | 3 | 02: 40; 04: 50; 09: < 24; 10: 87.07; 11: N.A.; 12: 190; 13: < 100 | 16
| 1 | 4 | 02: 40; 04: 70; 09: < 24; 10: 81.57; 11: N.A.; 12: 200; 13: < 100 | 94.1347
| 1 | 5 | 02: 40; 04: 50; 09: < 24; 10: 97.21; 11: N.A.; 12: 180; 13: < 100 | 8
| 1 | 6 | 02: 40; 04: 50; 09: < 24; 10: 97.6; 11: N.A.; 12: 200; 13: < 100
| Mean |  | 02: 38.3333; 04: 53.3333; 10: 91.5383; 12: 193.3333
| Median |  | 02: 40; 04: 50; 10: 92.14; 12: 195
| Std Dev. |  | 02: 4.0825; 04: 8.165; 10: 7.7605; 12: 8.165
| Rel.Std.Dev. |  | 02: 10.65%; 04: 15.31%; 10: 8.478%; 12: 4.223%
| PDM3 |  | 02: -59.28%; 04: -43.34%; 10: -2.758%; 12: 105.4%
| Z-Score (Absolute) |  | 02: 0.34; 04: 0.04; 09: 1.01; 10: 1.01; 11: NA; 12: 3.59; 13: 0.04
| Analytical results for P2O5 in OREAS 554 (Certified Value 0.073 wt.%) | [20 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 02 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 3
3
| 1 | 1 | 0.0733 | 0.071 | 0.07 | 0.073 | 0.0412 | 0.08 | 0.08 | 0.07 | 0.07 | 0.06 | 0.09 | 0.03952 | 0.07 | N.A. | 0.07966 | 0.07 | 0.067078 | 0.06 |  | 1
| 1 | 2 | 0.0733 | 0.073 | 0.07 | 0.075 | 0.0504 | 0.08 | 0.08 | 0.07 | 0.07 | 0.06 | 0.08 | 0.03694 | 0.07 | N.A. | 0.078345 | 0.072 | 0.06698 | 0.08 |  | 0
| 1 | 3 | 0.0733 | 0.073 | 0.08 | 0.073 | 0.0504 | 0.08 | 0.08 | 0.07 | 0.07 | 0.06 | 0.09 | 0.04271 | 0.07 | N.A. | 0.081696 | 0.072 | 0.066745 | 0.07 |  | 16
| 1 | 4 | 0.0733 | 0.072 | 0.08 | 0.075 | 0.0573 | 0.08 | 0.08 | 0.07 | 0.07 | 0.06 | 0.08 | 0.04255 | 0.06 | N.A. | 0.078234 | 0.072 | 0.067078 | 0.07 |  | 0.07316
| 1 | 5 | 0.0733 | 0.073 | 0.07 | 0.072 | 0.055 | 0.08 | 0.08 | 0.07 | 0.07 | 0.06 | 0.09 | 0.04163 | 0.07 | N.A. | 0.079634 | 0.073 | 0.067176 | 0.07 |  | 17
| 1 | 6 | 0.0733 | 0.071 | 0.07 | 0.073 | 0.0458 | 0.08 | 0.08 | 0.07 | 0.08 | 0.06 | 0.08 | 0.04207 | 0.07 | N.A. | 0.078885 | 0.073 | 0.066881 | 0.08
| Mean |  | 0.0733 | 0.072167 | 0.073333 | 0.0735 | 0.050017 | 0.08 | 0.08 | 0.07 | 0.071667 | 0.06 | 0.085 | 0.040903 | 0.068333 |  | 0.079409 | 0.072 | 0.06699 | 0.071667
| Median |  | 0.0733 | 0.0725 | 0.07 | 0.073 | 0.0504 | 0.08 | 0.08 | 0.07 | 0.07 | 0.06 | 0.085 | 0.04185 | 0.07 |  | 0.07926 | 0.072 | 0.067029 | 0.07
| Std Dev. |  | 1.5202e-17 | 0.00098319 | 0.005164 | 0.0012247 | 0.0058939 | 0 | 0 | 0 | 0.0040825 | 0 | 0.0054772 | 0.0022577 | 0.0040825 |  | 0.0012755 | 0.0010954 | 0.00015624 | 0.0075277
| Rel.Std.Dev. |  | 2.074e-14% | 1.362% | 7.042% | 1.666% | 11.78% | 0% | 0% | 0% | 5.696% | 0% | 6.444% | 5.52% | 5.974% |  | 1.606% | 1.521% | 0.2332% | 10.5%
| PDM3 |  | 0.1917% | -1.357% | 0.2372% | 0.465% | -31.63% | 9.35% | 9.35% | -4.319% | -2.041% | -17.99% | 16.18% | -44.09% | -6.597% |  | 8.542% | -1.585% | -8.434% | -2.041%
| Z-Score (Absolute) |  | 0.24 | 0.03 | 0.25 | 0.28 | 4.04 | 1.47 | 1.47 | 0.37 | 0.06 | 2.21 | 2.39 | 5.72 | 0.67 | NA | 1.36 | 0 | 0.92 | 0.06
| Analytical results for Pb in OREAS 554 (Indicative Value 86 ppm) | [20 empty] | Indicative
Pb | Round | Replicate | 02: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 02: 02; 09: 12; 10: 19; 11: 21 |  | ppm
|  |  | 02: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF |  | 0
0
| 1 | 1 | 02: < 10; 09: 20; 10: < 90; 11: 100 |  | 1
| 1 | 2 | 02: < 10; 09: 30; 10: < 90; 11: < 100 |  | 0
| 1 | 3 | 02: < 10; 09: 30; 10: < 90; 11: 200 |  | 16
| 1 | 4 | 02: < 10; 09: 40; 10: < 90; 11: 200 |  | 85.8333
| 1 | 5 | 02: < 10; 09: 60; 10: < 90; 11: 100 |  | 9
| 1 | 6 | 02: < 10; 09: 30; 10: < 90; 11: 200
| Mean |  | 09: 35; 11: 160
| Median |  | 09: 30; 11: 200
| Std Dev. |  | 09: 13.784; 11: 54.7723
| Rel.Std.Dev. |  | 09: 39.38%; 11: 34.23%
| PDM3 |  | 09: -59.22%; 11: 86.41%
| Z-Score (Absolute) |  | 02: 1.16; 09: 0.19; 10: 0.19; 11: 3.29
| Analytical results for Sb in OREAS 554 (Indicative Value < 10 ppm) | [20 empty] | Indicative
Sb | Round | Replicate | 02: Lab |  | 1
| No | No | 02: 12 |  | ppm
|  |  | 02: BF*XRF |  | 1
1
| 1 | 1 | 02: < 10 |  | 1
| 1 | 2 | 02: < 10 |  | 0
| 1 | 3 | 02: < 10 |  | 16
| 1 | 4 | 02: < 10 |  | < 10
| 1 | 5 | 02: < 10 |  | 10
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for SiO2 in OREAS 554 (Certified Value 60.26 wt.%) | [20 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 60.57 | 61.6 | 60.89 | 60.9 | 60.8 | 60.197 | 60.24 | 60.6 | 59.92 | 60.65 | 59.98 | 60.5 | 60.8214 | 59.4 | 61.1 | 60.1506 | 59.99 | 61.745 | 60.06 | 1
| 1 | 2 | 60.77 | 61.49 | 60.89 | 60.07 | 60.6 | 60.197 | 60.39 | 60.12 | 60.24 | 60.55 | 60.13 | 59.5 | 61.1144 | 60.2 | 58.8 | 59.8444 | 59.87 | 61.6321 | 59.98 | 0
| 1 | 3 | 60.43 | 61.69 | 60.81 | 60.42 | 60.7 | 60.283 | 60.24 | 60.64 | 60.15 | 60.33 | 59.74 | 59.8 | 60.8308 | 59.9 | 60.2 | 59.6912 | 60.2 | 61.6084 | 60.01 | 16
| 1 | 4 | 60.52 | 61.53 | 60.91 | 59.91 | 60.7 | 59.021 | 59.96 | 60.62 | 60.36 | 60.21 | 60.31 | 60 | 60.6106 | 59.1 | 62.3 | 59.8007 | 60.47 | 61.7555 | 60.07 | 60.2623
| 1 | 5 | 60.56 | 61.46 | 60.99 | 60.42 | 60.6 | 61.053 | 59.96 | 60.81 | 60.66 | 60.48 | 61.04 | 60 | 60.9021 | 59.7 | 60.6 | 59.7217 | 60.24 | 61.6268 | 59.93 | 18
| 1 | 6 | 60.41 | 61.62 | 60.88 | 60.34 | 60.8 | 59.577 | 59.89 | 60.14 | 59.83 | 60.7 | 59.48 | 60 | 60.9733 | 59.3 | 59.8 | 59.8932 | 60.02 | 61.649 | 59.97
| Mean |  | 60.5433 | 61.565 | 60.895 | 60.3433 | 60.7 | 60.0547 | 60.1133 | 60.4883 | 60.1933 | 60.4867 | 60.1133 | 59.9667 | 60.8754 | 59.6 | 60.4667 | 59.8503 | 60.1317 | 61.6695 | 60.0033
| Median |  | 60.54 | 61.565 | 60.89 | 60.38 | 60.7 | 60.197 | 60.1 | 60.61 | 60.195 | 60.515 | 60.055 | 60 | 60.8665 | 59.55 | 60.4 | 59.8226 | 60.11 | 61.6406 | 59.995
| Std Dev. |  | 0.12925 | 0.086891 | 0.057879 | 0.3418 | 0.089443 | 0.69079 | 0.20275 | 0.2875 | 0.3021 | 0.18833 | 0.53962 | 0.3266 | 0.16876 | 0.40988 | 1.1894 | 0.16515 | 0.21535 | 0.063979 | 0.054283
| Rel.Std.Dev. |  | 0.2135% | 0.1411% | 0.09505% | 0.5664% | 0.1474% | 1.15% | 0.3373% | 0.4753% | 0.5019% | 0.3114% | 0.8977% | 0.5446% | 0.2772% | 0.6877% | 1.967% | 0.2759% | 0.3581% | 0.1037% | 0.09047%
| PDM3 |  | 0.4664% | 2.162% | 1.05% | 0.1345% | 0.7264% | -0.3445% | -0.2471% | 0.3751% | -0.1144% | 0.3724% | -0.2471% | -0.4905% | 1.018% | -1.099% | 0.3392% | -0.6836% | -0.2167% | 2.335% | -0.4297%
| Z-Score (Absolute) |  | 0.47 | 2.85 | 1.29 | 0 | 0.83 | 0.67 | 0.54 | 0.34 | 0.35 | 0.33 | 0.54 | 0.88 | 1.24 | 1.74 | 0.29 | 1.15 | 0.49 | 3.1 | 0.79
| Analytical results for Sn in OREAS 554 (Indicative Value < 10 ppm) | [20 empty] | Indicative
Sn | Round | Replicate | 02: Lab | 1
| No | No | 02: 02 | ppm
|  |  | 02: BF*XRF | 1
1
| 1 | 1 | 02: < 10 | 1
| 1 | 2 | 02: < 10 | 0
| 1 | 3 | 02: < 10 | 16
| 1 | 4 | 02: < 10 | < 10
| 1 | 5 | 02: < 10 | 11
| 1 | 6 | 02: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 02
| Analytical results for SO3 in OREAS 554 (Certified Value 1.98 wt.%) | [20 empty] | Certified
SO3 | Round | Replicate | 02: Lab; 04: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab; 17: Lab | 1
| No | No | 02: 02; 04: 12; 09: 13; 10: 14; 11: 15; 12: 16; 13: 17; 14: 19; 15: 21; 16: 23; 17: 25 | wt.%
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF; 12: BF*XRF; 13: BF*XRF; 14: BF*XRF; 15: BF*XRF; 16: BF*XRF; 17: BF*XRF | 2
3
| 1 | 1 | 02: 2.0377; 04: 2.3224; 09: 1.97; 10: 1.97; 11: 1.97; 12: 2; 13: 2.01; 14: 1.9354; 15: 1.973; 16: 1.918; 17: 2.0002 | 1
| 1 | 2 | 02: 2.0302; 04: 2.2475; 09: 2.02; 10: 1.95; 11: 1.95; 12: 2; 13: 2.01; 14: 1.9241; 15: 1.748; 16: 1.911; 17: 1.9982 | 0
| 1 | 3 | 02: 2.0152; 04: 2.24; 09: 2; 10: 1.98; 11: 1.94; 12: 1.97; 13: 1.98; 14: 1.8691; 15: 1.798; 16: 1.817; 17: 2.0053 | 16
| 1 | 4 | 02: 2.0202; 04: 2.2225; 09: 2; 10: 1.96; 11: 1.95; 12: 1.98; 13: 2.01; 14: 1.8802; 15: 1.848; 16: 1.938; 17: 2.0088 | 1.9838
| 1 | 5 | 02: 2.0227; 04: 2.2974; 09: 1.95; 10: 1.96; 11: 2.03; 12: 1.98; 13: 2; 14: 1.8511; 15: 1.873; 16: 1.956; 17: 1.9968 | 19
| 1 | 6 | 02: 2.0427; 04: 2.3474; 09: 1.97; 10: 1.96; 11: 1.98; 12: 2; 13: 1.98; 14: 1.8897; 15: 1.798; 16: 1.951; 17: 2.0043
| Mean |  | 02: 2.0281; 04: 2.2795; 09: 1.985; 10: 1.9633; 11: 1.97; 12: 1.9883; 13: 1.9983; 14: 1.8916; 15: 1.8397; 16: 1.9152; 17: 2.0023
| Median |  | 02: 2.0265; 04: 2.2725; 09: 1.985; 10: 1.96; 11: 1.96; 12: 1.99; 13: 2.005; 14: 1.885; 15: 1.823; 16: 1.928; 17: 2.0023
| Std Dev. |  | 02: 0.010656; 04: 0.050209; 09: 0.025884; 10: 0.010328; 11: 0.032863; 12: 0.013292; 13: 0.01472; 14: 0.03245; 15: 0.078528; 16: 0.051254; 17: 0.0046073
| Rel.Std.Dev. |  | 02: 0.5254%; 04: 2.203%; 09: 1.304%; 10: 0.526%; 11: 1.668%; 12: 0.6685%; 13: 0.7366%; 14: 1.715%; 15: 4.269%; 16: 2.676%; 17: 0.2301%
| PDM3 |  | 02: 2.235%; 04: 14.91%; 09: 0.06169%; 10: -1.031%; 11: -0.6944%; 12: 0.2297%; 13: 0.7338%; 14: -4.646%; 15: -7.264%; 16: -3.459%; 17: 0.9325%
| Z-Score (Absolute) |  | 02: 1.34; 04: 9.17; 09: 0; 10: 0.67; 11: 0.47; 12: 0.1; 13: 0.41; 14: 2.91; 15: 4.53; 16: 2.17; 17: 0.54
| Analytical results for SrO in OREAS 554 (Certified Value 94 ppm) | [20 empty] | Certified
SrO | Round | Replicate | 02: Lab; 04: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 02: 02; 04: 13; 09: 14; 10: 15; 11: 16; 12: 17; 13: 19; 14: 20; 15: 28 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF; 12: BF*XRF; 13: BF*XRF; 14: BF*XRF; 15: BF*XRF | 0
0
| 1 | 1 | 02: 83; 04: 100; 09: 100; 10: 100; 11: 100; 12: 100; 13: < 100; 14: < 100; 15: < 100 | 1
| 1 | 2 | 02: 83; 04: 100; 09: 100; 10: 100; 11: 100; 12: 100; 13: < 100; 14: < 100; 15: < 100 | 0
| 1 | 3 | 02: 71; 04: 100; 09: 100; 10: 100; 11: 100; 12: 100; 13: < 100; 14: < 100; 15: < 100 | 16
| 1 | 4 | 02: 83; 04: 100; 09: 100; 10: 100; 11: 100; 12: 100; 13: < 100; 14: < 100; 15: < 100 | 94.0241
| 1 | 5 | 02: 83; 04: 100; 09: 100; 10: 200; 11: 100; 12: < 100; 13: < 100; 14: < 100; 15: < 100 | 20
| 1 | 6 | 02: 83; 04: 100; 09: 100; 10: 100; 11: 100; 12: < 100; 13: < 100; 14: < 100; 15: < 100
| Mean |  | 02: 81; 04: 100; 09: 100; 10: 116.6667; 11: 100; 12: 100
| Median |  | 02: 83; 04: 100; 09: 100; 10: 100; 11: 100; 12: 100
| Std Dev. |  | 02: 4.899; 04: 0; 09: 0; 10: 40.8248; 11: 0; 12: 0
| Rel.Std.Dev. |  | 02: 6.048%; 04: 0%; 09: 0%; 10: 34.99%; 11: 0%; 12: 0%
| PDM3 |  | 02: -13.85%; 04: 6.356%; 09: 6.356%; 10: 24.08%; 11: 6.356%; 12: 6.356%
| Z-Score (Absolute) |  | 02: 0.1; 04: 0.67; 09: 0.67; 10: 1.35; 11: 0.67; 12: 0; 13: 1.35; 14: 1.35; 15: 1.35
| Analytical results for TiO2 in OREAS 554 (Certified Value 0.382 wt.%) | [20 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 25 | 28 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.38 | 0.39 | 0.39 | 0.38 | 0.39 | 0.3503 | 0.38 | 0.4 | 0.37 | 0.38 | 0.41 | 0.47 | 0.41782 | 0.37 | 0.37 | 0.37854 | 0.39 | 0.38376 | 0.37 | 1
| 1 | 2 | 0.38 | 0.38 | 0.39 | 0.37 | 0.38 | 0.3837 | 0.38 | 0.4 | 0.37 | 0.38 | 0.41 | 0.44 | 0.40988 | 0.37 | 0.36 | 0.3875 | 0.38 | 0.38475 | 0.37 | 0
| 1 | 3 | 0.38 | 0.38 | 0.39 | 0.38 | 0.39 | 0.3753 | 0.39 | 0.4 | 0.37 | 0.38 | 0.39 | 0.44 | 0.40719 | 0.38 | 0.41 | 0.3866 | 0.39 | 0.38455 | 0.36 | 16
| 1 | 4 | 0.38 | 0.39 | 0.39 | 0.36 | 0.38 | 0.367 | 0.39 | 0.41 | 0.37 | 0.38 | 0.39 | 0.46 | 0.4308 | 0.38 | 0.38 | 0.39395 | 0.39 | 0.38711 | 0.37 | 0.3818
| 1 | 5 | 0.38 | 0.38 | 0.39 | 0.36 | 0.38 | 0.3586 | 0.39 | 0.4 | 0.37 | 0.38 | 0.38 | 0.46 | 0.40686 | 0.38 | 0.4 | 0.39378 | 0.38 | 0.38494 | 0.38 | 21
| 1 | 6 | 0.38 | 0.38 | 0.39 | 0.38 | 0.38 | 0.367 | 0.37 | 0.4 | 0.37 | 0.38 | 0.38 | 0.47 | 0.40979 | 0.38 | 0.38 | 0.39048 | 0.38 | 0.38573 | 0.36
| Mean |  | 0.38 | 0.38333 | 0.39 | 0.37167 | 0.38333 | 0.36698 | 0.38333 | 0.40167 | 0.37 | 0.38 | 0.39333 | 0.45667 | 0.41372 | 0.37667 | 0.38333 | 0.38848 | 0.385 | 0.38514 | 0.36833
| Median |  | 0.38 | 0.38 | 0.39 | 0.375 | 0.38 | 0.367 | 0.385 | 0.4 | 0.37 | 0.38 | 0.39 | 0.46 | 0.40984 | 0.38 | 0.38 | 0.38899 | 0.385 | 0.38484 | 0.37
| Std Dev. |  | 6.0809e-17 | 0.005164 | 6.0809e-17 | 0.0098319 | 0.005164 | 0.011809 | 0.008165 | 0.0040825 | 6.0809e-17 | 6.0809e-17 | 0.013663 | 0.013663 | 0.0092566 | 0.005164 | 0.018619 | 0.0057478 | 0.0054772 | 0.0011528 | 0.0075277
| Rel.Std.Dev. |  | 1.6e-14% | 1.347% | 1.559e-14% | 2.645% | 1.347% | 3.218% | 2.13% | 1.016% | 1.643e-14% | 1.6e-14% | 3.474% | 2.992% | 2.237% | 1.371% | 4.857% | 1.48% | 1.423% | 0.2993% | 2.044%
| PDM3 |  | -0.4714% | 0.4016% | 2.148% | -2.654% | 0.4016% | -3.881% | 0.4016% | 5.203% | -3.091% | -0.4714% | 3.021% | 19.61% | 8.361% | -1.344% | 0.4016% | 1.749% | 0.8382% | 0.8748% | -3.527%
| Z-Score (Absolute) |  | 0.34 | 0 | 0.67 | 1.18 | 0 | 1.65 | 0 | 1.85 | 1.35 | 0.34 | 1.01 | 7.42 | 3.07 | 0.67 | 0 | 0.52 | 0.17 | 0.18 | 1.52
| Analytical results for V2O5 in OREAS 554 (Indicative Value 208 ppm) | [20 empty] | Indicative
V2O5 | Round | Replicate | 02: Lab; 04: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 02: 02; 04: 10; 09: 11; 10: 19; 11: 20 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF | 0
0
| 1 | 1 | 02: 210; 04: 250; 09: 100; 10: 215.5; 11: 200 | 1
| 1 | 2 | 02: 220; 04: 260; 09: 100; 10: 179; 11: 200 | 0
| 1 | 3 | 02: 220; 04: 250; 09: 200; 10: 193.3; 11: 200 | 16
| 1 | 4 | 02: 210; 04: 260; 09: 100; 10: 222; 11: 200 | 207.6444
| 1 | 5 | 02: 230; 04: 250; 09: 100; 10: 216.6; 11: 200 | 7
| 1 | 6 | 02: 210; 04: 250; 09: 100; 10: 211.2; 11: 200
| Mean |  | 02: 216.6667; 04: 253.3333; 09: 116.6667; 10: 206.2667; 11: 200
| Median |  | 02: 215; 04: 250; 09: 100; 10: 213.35; 11: 200
| Std Dev. |  | 02: 8.165; 04: 5.164; 09: 40.8248; 10: 16.5862; 11: 0
| Rel.Std.Dev. |  | 02: 3.768%; 04: 2.038%; 09: 34.99%; 10: 8.041%; 11: 0%
| PDM3 |  | 02: 4.345%; 04: 22%; 09: -43.81%; 10: -0.6635%; 11: -3.682%
| Z-Score (Absolute) |  | 02: 0.67; 04: 3.05; 09: 5.81; 10: 0; 11: 0.41
| Analytical results for Zn in OREAS 554 (Indicative Value 66 ppm) | [20 empty] | Indicative
Zn | Round | Replicate | 02: Lab; 04: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab | 1
| No | No | 02: 02; 04: 09; 09: 12; 10: 19; 11: 21; 12: 23; 13: 25 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF; 09: BF*XRF; 10: BF*XRF; 11: BF*XRF; 12: BF*XRF; 13: BF*XRF | 0
0
| 1 | 1 | 02: 60; 04: 50; 09: 340; 10: < 80; 11: 100; 12: < 50; 13: < 100 | 1
| 1 | 2 | 02: 50; 04: 50; 09: 430; 10: < 80; 11: 100; 12: < 50; 13: < 100 | 0
| 1 | 3 | 02: 60; 04: 40; 09: 420; 10: < 80; 11: 100; 12: < 50; 13: < 100 | 16
| 1 | 4 | 02: 70; 04: 50; 09: 440; 10: < 80; 11: 100; 12: < 50; 13: < 100 | 66.1111
| 1 | 5 | 02: 60; 04: 50; 09: 560; 10: < 80; 11: 100; 12: < 50; 13: < 100 | 8
| 1 | 6 | 02: 60; 04: 40; 09: 430; 10: < 80; 11: < 100; 12: < 50; 13: < 100
| Mean |  | 02: 60; 04: 46.6667; 09: 436.6667; 11: 100
| Median |  | 02: 60; 04: 50; 09: 430; 11: 100
| Std Dev. |  | 02: 6.3246; 04: 5.164; 09: 70.6163; 11: 0
| Rel.Std.Dev. |  | 02: 10.54%; 04: 11.07%; 09: 16.17%; 11: 0%
| PDM3 |  | 02: -9.244%; 04: -29.41%; 09: 560.5%; 11: 51.26%
| Z-Score (Absolute) |  | 02: 0.67; 04: 0.22; 09: 26.07; 10: 0.66; 11: 2.81; 12: 1.69; 13: 0
| Analytical results for Zr in OREAS 554 (Indicative Value 67 ppm) | [20 empty] | Indicative
Zr | Round | Replicate | 02: Lab; 04: Lab | 1
| No | No | 02: 02; 04: 19 | ppm
|  |  | 02: BF*XRF; 04: BF*XRF | 0
0
| 1 | 1 | 02: 70; 04: < 70 | 1
| 1 | 2 | 02: 70; 04: < 70 | 0
| 1 | 3 | 02: 60; 04: < 70 | 16
| 1 | 4 | 02: 70; 04: < 70 | 66.6667
| 1 | 5 | 02: 70; 04: < 70 | 9
| 1 | 6 | 02: 60; 04: < 70
| Mean |  | 02: 66.6667
| Median |  | 02: 70
| Std Dev. |  | 02: 5.164
| Rel.Std.Dev. |  | 02: 7.746%
| PDM3 |  | 02: -4.441e-14%
| Z-Score (Absolute) |  | 02: 0.67; 04: 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 554 (Indicative Value 0.128 wt.%) | [25 empty] | Indicative
H2O- | Round | Replicate | Lab | [23 empty] | 1
| No | No | 19 | [23 empty] | wt.%
|  |  | BF*XRF | [23 empty] | 3
3
| 1 | 1 | 0.1 | [23 empty] | 1
| 1 | 2 | 0.11 | [23 empty] | 0
| 1 | 3 | 0.12 | [23 empty] | 16
| 1 | 4 | 0.12 | [23 empty] | 0.12833
| 1 | 5 | 0.15 | [23 empty] | 13
| 1 | 6 | 0.17
| Mean |  | 0.12833
| Median |  | 0.12
| Std Dev. |  | 0.026394
| Rel.Std.Dev. |  | 20.57%
| PDM3 |  | 2.665e-13%
| Z-Score (Absolute) |  | NA
Note: some laboratories include H2O- in their fusion package.
| Analytical results for LOI1000 in OREAS 554 (Certified Value 10.11 wt.%) | [25 empty] | Certified
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 28 | 29 | 30 | wt.%
|  |  | LOI*TGA | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | BF*XRF | LOI*Furnace | BF*XRF | LOI*Furnace | LOI*TGA | LOI*Furnace | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | BF*XRF | LOI*Furnace | BF*XRF | LOI*TGA | LOI*TGA | LOI*TGA | BF*XRF | LOI*TGA | LOI*TGA | 2
3
| 1 | 1 | 9.89 | 10.01 | 9.97 | 10.04 | 10.2 | 10.46 | 9.94 | 9.6 | 9.83 | 9.96 | 10.1 | 9.94 | 10.38 | 10.72 | 9.99 | 10.2 | 10.61 | 10.1912 | 10.004 | 9.93 | 10.2671 | 10.15 | 10.34 | 11.4 | 1
| 1 | 2 | 9.91 | 10.01 | 9.94 | 10.02 | 10.3 | 10.4 | 9.95 | 9.53 | 9.9 | 9.96 | 10.11 | 9.96 | 10.38 | 10.7 | 10 | 10.2 | 10.4 | 10.2583 | 10.12 | 10.02 | 10.2582 | 10.17 | 10.22 | 11.5 | 0
| 1 | 3 | 9.91 | 9.98 | 9.95 | 14.07 | 10.4 | 10.45 | 9.92 | 9.47 | 9.96 | 9.95 | 10.01 | 9.91 | 10.2 | 10.7 | 10.09 | 10.2 | 10.36 | 10.3222 | 10.158 | 10.06 | 10.3846 | 10.18 | 10.26 | 11.4 | 16
| 1 | 4 | 9.86 | 9.98 | 9.98 | 9.82 | 10.5 | 10.44 | 9.96 | 9.45 | 9.93 | 9.95 | 10.06 | 9.92 | 10.17 | 10.69 | 10.12 | 10.2 | 10.5 | 10.2819 | 10.147 | 10.13 | 10.1546 | 10.14 | 10.31 | 11.5 | 10.1141
| 1 | 5 | 9.89 | 9.99 | 9.93 | 12.17 | 10.3 | 10.38 | 9.95 | 9.47 | 9.99 | 9.94 | 10.12 | 9.91 | 10.12 | 10.63 | 10.05 | 10.2 | 10.3 | 10.3219 | 10.074 | 10.2 | 10.2306 | 10.12 | 10.35 | 11.6 | 23
| 1 | 6 | 9.92 | 9.97 | 9.94 | 9.82 | 10.5 | 10.48 | 9.88 | 9.52 | 9.83 | 9.94 | 10.06 | 9.94 | 10.14 | 10.62 | 10.5 | 10.2 | 10.64 | 10.2561 | 10.127 | 10.16 | 10.2202 | 10.03 | 10.23 | 11.6
| Mean |  | 9.8967 | 9.99 | 9.9517 | 10.99 | 10.3667 | 10.435 | 9.9333 | 9.5067 | 9.9067 | 9.95 | 10.0767 | 9.93 | 10.2317 | 10.6767 | 10.125 | 10.2 | 10.4683 | 10.2719 | 10.105 | 10.0833 | 10.2526 | 10.1317 | 10.285 | 11.5
| Median |  | 9.9 | 9.985 | 9.945 | 10.03 | 10.35 | 10.445 | 9.945 | 9.495 | 9.915 | 9.95 | 10.08 | 9.93 | 10.185 | 10.695 | 10.07 | 10.2 | 10.45 | 10.2701 | 10.1235 | 10.095 | 10.2444 | 10.145 | 10.285 | 11.5
| Std Dev. |  | 0.021602 | 0.016733 | 0.019408 | 1.7584 | 0.12111 | 0.037815 | 0.029439 | 0.055377 | 0.066533 | 0.0089443 | 0.041312 | 0.02 | 0.11805 | 0.041312 | 0.19045 | 1.9459e-15 | 0.13805 | 0.049147 | 0.057348 | 0.099733 | 0.075878 | 0.054191 | 0.056125 | 0.089443
| Rel.Std.Dev. |  | 0.2183% | 0.1675% | 0.195% | 16% | 1.168% | 0.3624% | 0.2964% | 0.5825% | 0.6716% | 0.08989% | 0.41% | 0.2014% | 1.154% | 0.3869% | 1.881% | 1.908e-14% | 1.319% | 0.4785% | 0.5675% | 0.9891% | 0.7401% | 0.5349% | 0.5457% | 0.7778%
| PDM3 |  | -2.15% | -1.227% | -1.606% | 8.66% | 2.497% | 3.173% | -1.787% | -6.006% | -2.051% | -1.623% | -0.3702% | -1.82% | 1.162% | 5.562% | 0.1077% | 0.8492% | 3.502% | 1.56% | -0.09005% | -0.3043% | 1.369% | 0.1736% | 1.69% | 13.7%
| Z-Score (Absolute) |  | 0.88 | 0.53 | 0.67 | 3.27 | 0.91 | 1.17 | 0.74 | 2.36 | 0.84 | 0.68 | 0.2 | 0.75 | 0.39 | 2.08 | 0.01 | 0.27 | 1.29 | 0.55 | 0.09 | 0.17 | 0.47 | 0.01 | 0.6 | 5.21
Note: some laboratories include LOI in their fusion package.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 554 (Certified Value 2.84 wt.%) | [28 empty] | Certified
C | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 28 | 29 | 30 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 2.737 | 2.91 | 2.84 | 3 | 2.8 | 2.64 | 2.705 | 2.83 | 2.85 | 2.86 | 2.93 | 2.92 | 2.9 | 2.77 | 2.91 | 2.79 | 2.8 | 2.83 | 2.82 | 2.92 | 2.8037 | 2.67 | 2.85 | 2.7892 | 2.78 | 2.8 | 2.86 | 1
| 1 | 2 | 2.876 | 2.91 | 2.84 | 3 | 2.9 | 2.62 | 2.6753 | 2.86 | 2.85 | 2.8 | 2.92 | 2.93 | 2.9 | 2.77 | 2.96 | 2.79 | 1.62 | 2.805 | 2.81 | 2.93 | 2.7865 | 2.67 | 2.87 | 2.789 | 2.78 | 2.84 | 2.84 | 0
| 1 | 3 | 2.721 | 2.84 | 2.83 | 3.02 | 2.8 | 2.66 | 2.657 | 2.79 | 2.87 | 2.81 | 2.9 | 2.93 | 2.92 | 2.79 | 2.93 | 2.8 | 2.84 | 2.84 | 2.82 | 2.91 | 2.8305 | 2.67 | 2.86 | 2.797 | 2.77 | 2.83 | 2.86 | 16
| 1 | 4 | 2.893 | 2.85 | 2.82 | 2.96 | 2.8 | 2.66 | 2.454 | 2.87 | 2.85 | 2.83 | 2.97 | 2.94 | 2.9 | 2.76 | 2.84 | 2.79 | 2.82 | 2.83 | 2.83 | 2.76 | 2.8111 | 2.67 | 2.85 | 2.799 | 2.79 | 2.8 | 2.83 | 2.8436
| 1 | 5 | 2.708 | 2.91 | 2.83 | 3.02 | 2.9 | 2.67 | 2.556 | 2.86 | 2.87 | 2.81 | 2.96 | 2.93 | 2.93 | 2.73 | 2.83 | 2.81 | 2.78 | 2.825 | 2.82 | 2.86 | 2.8076 | 2.68 | 2.91 | 2.7922 | 2.78 | 2.83 | 2.84 | 25
| 1 | 6 | 2.916 | 2.91 | 2.89 | 3.04 | 2.9 | 2.62 | 2.582 | 2.84 | 2.87 | 2.81 | 2.93 | 2.92 | 2.9 | 2.72 | 2.85 | 2.83 | 2.83 | 2.815 | 2.84 | 2.94 | 2.7513 | 2.67 | 2.86 | 2.783 | 2.77 | 2.81 | 2.86
| Mean |  | 2.8085 | 2.8883 | 2.8417 | 3.0067 | 2.85 | 2.645 | 2.6049 | 2.8417 | 2.86 | 2.82 | 2.935 | 2.9283 | 2.9083 | 2.7567 | 2.8867 | 2.8017 | 2.615 | 2.8242 | 2.8233 | 2.8867 | 2.7984 | 2.6717 | 2.8667 | 2.7916 | 2.7783 | 2.8183 | 2.8483
| Median |  | 2.8065 | 2.91 | 2.835 | 3.01 | 2.85 | 2.65 | 2.6195 | 2.85 | 2.86 | 2.81 | 2.93 | 2.93 | 2.9 | 2.765 | 2.88 | 2.795 | 2.81 | 2.8275 | 2.82 | 2.915 | 2.8056 | 2.67 | 2.86 | 2.7907 | 2.78 | 2.82 | 2.85
| Std Dev. |  | 0.096043 | 0.033714 | 0.024833 | 0.027325 | 0.054772 | 0.021679 | 0.093147 | 0.029269 | 0.010954 | 0.021909 | 0.025884 | 0.0075277 | 0.013292 | 0.026583 | 0.053914 | 0.016021 | 0.48792 | 0.012416 | 0.010328 | 0.06802 | 0.027072 | 0.0040825 | 0.022509 | 0.0058436 | 0.0075277 | 0.017224 | 0.013292
| Rel.Std.Dev. |  | 3.42% | 1.167% | 0.8739% | 0.9088% | 1.922% | 0.8196% | 3.576% | 1.03% | 0.383% | 0.7769% | 0.8819% | 0.2571% | 0.457% | 0.9643% | 1.868% | 0.5718% | 18.66% | 0.4396% | 0.3658% | 2.356% | 0.9674% | 0.1528% | 0.7852% | 0.2093% | 0.2709% | 0.6111% | 0.4666%
| PDM3 |  | -1.233% | 1.574% | -0.06679% | 5.736% | 0.2263% | -6.983% | -8.394% | -0.06679% | 0.5779% | -0.8287% | 3.215% | 2.981% | 2.278% | -3.056% | 1.516% | -1.473% | -8.038% | -0.6822% | -0.7115% | 1.516% | -1.587% | -6.045% | 0.8124% | -1.829% | -2.294% | -0.8874% | 0.1677%
| Z-Score (Absolute) |  | 0.25 | 1.02 | 0.28 | 2.9 | 0.41 | 2.84 | 3.48 | 0.28 | 0.57 | 0.07 | 1.76 | 1.65 | 1.34 | 1.07 | 0.99 | 0.36 | 3.32 | 0 | 0.01 | 0.99 | 0.41 | 2.42 | 0.67 | 0.52 | 0.73 | 0.09 | 0.38
| Analytical results for S in OREAS 554 (Certified Value 0.771 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 28 | 29 | 30 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 3
3
| 1 | 1 | 0.733 | 0.78 | 0.75 | 0.78 | 0.7 | 0.81 | 0.85509 | 0.86 | 0.78 | 0.78 | 0.78 | 0.76 | 0.81 | 0.77 | 0.78 | 0.74 | 0.75 | 0.767 | 0.791 | 0.81 | 0.862 | 0.77 | 0.78 | 0.8005 | 0.76 | 0.79 | 0.75 | 1
| 1 | 2 | 0.729 | 0.8 | 0.75 | 0.77 | 0.8 | 0.84 | 0.793 | 0.87 | 0.77 | 0.77 | 0.78 | 0.78 | 0.81 | 0.76 | 0.76 | 0.78 | 0.58 | 0.7635 | 0.798 | 0.84 | 0.841 | 0.76 | 0.79 | 0.80435 | 0.77 | 0.78 | 0.75 | 0
| 1 | 3 | 0.745 | 0.73 | 0.72 | 0.77 | 0.7 | 0.84 | 0.842 | 0.86 | 0.77 | 0.77 | 0.79 | 0.76 | 0.8 | 0.77 | 0.76 | 0.77 | 0.78 | 0.788 | 0.802 | 0.83 | 0.87 | 0.77 | 0.77 | 0.80558 | 0.77 | 0.79 | 0.74 | 16
| 1 | 4 | 0.739 | 0.77 | 0.8 | 0.77 | 0.8 | 0.86 | 0.869 | 0.86 | 0.76 | 0.77 | 0.77 | 0.76 | 0.79 | 0.76 | 0.76 | 0.77 | 0.76 | 0.752 | 0.817 | 0.8 | 0.898 | 0.76 | 0.8 | 0.804 | 0.78 | 0.78 | 0.75 | 0.77127
| 1 | 5 | 0.744 | 0.77 | 0.77 | 0.77 | 0.7 | 0.84 | 0.81 | 0.88 | 0.77 | 0.75 | 0.77 | 0.77 | 0.79 | 0.75 | 0.75 | 0.76 | 0.76 | 0.7615 | 0.804 | 0.84 | 0.892 | 0.76 | 0.79 | 0.80324 | 0.79 | 0.79 | 0.74 | 26
| 1 | 6 | 0.749 | 0.79 | 0.71 | 0.77 | 0.6 | 0.83 | 0.82933 | 0.86 | 0.77 | 0.77 | 0.78 | 0.77 | 0.79 | 0.75 | 0.77 | 0.75 | 0.76 | 0.7835 | 0.808 | 0.82 | 0.781 | 0.76 | 0.79 | 0.806 | 0.77 | 0.8 | 0.74
| Mean |  | 0.73983 | 0.77333 | 0.75 | 0.77167 | 0.71667 | 0.83667 | 0.83307 | 0.865 | 0.77 | 0.76833 | 0.77833 | 0.76667 | 0.79833 | 0.76 | 0.76333 | 0.76167 | 0.73167 | 0.76925 | 0.80333 | 0.82333 | 0.85733 | 0.76333 | 0.78667 | 0.80394 | 0.77333 | 0.78833 | 0.745
| Median |  | 0.7415 | 0.775 | 0.75 | 0.77 | 0.7 | 0.84 | 0.83567 | 0.86 | 0.77 | 0.77 | 0.78 | 0.765 | 0.795 | 0.76 | 0.76 | 0.765 | 0.76 | 0.76525 | 0.803 | 0.825 | 0.866 | 0.76 | 0.79 | 0.80417 | 0.77 | 0.79 | 0.745
| Std Dev. |  | 0.0076529 | 0.024221 | 0.032863 | 0.0040825 | 0.075277 | 0.01633 | 0.02831 | 0.0083666 | 0.0063246 | 0.0098319 | 0.0075277 | 0.008165 | 0.0098319 | 0.0089443 | 0.010328 | 0.01472 | 0.074944 | 0.013787 | 0.0088468 | 0.01633 | 0.042735 | 0.005164 | 0.010328 | 0.0019718 | 0.010328 | 0.0075277 | 0.0054772
| Rel.Std.Dev. |  | 1.034% | 3.132% | 4.382% | 0.529% | 10.5% | 1.952% | 3.398% | 0.9672% | 0.8214% | 1.28% | 0.9672% | 1.065% | 1.232% | 1.177% | 1.353% | 1.933% | 10.24% | 1.792% | 1.101% | 1.983% | 4.985% | 0.6765% | 1.313% | 0.2453% | 1.336% | 0.9549% | 0.7352%
| PDM3 |  | -4.076% | 0.2674% | -2.758% | 0.05128% | -7.08% | 8.479% | 8.013% | 12.15% | -0.1648% | -0.3809% | 0.9157% | -0.597% | 3.509% | -1.461% | -1.029% | -1.245% | -5.135% | -0.2621% | 4.157% | 6.75% | 11.16% | -1.029% | 1.996% | 4.236% | 0.2674% | 2.212% | -3.406%
| Z-Score (Absolute) |  | 1.47 | 0.04 | 1.01 | 0.04 | NA | 2.88 | 2.72 | 4.16 | 0.11 | 0.19 | 0.26 | 0.26 | 1.16 | 0.56 | 0.41 | 0.49 | 1.84 | 0.15 | 1.39 | 2.29 | 3.81 | 0.41 | 0.64 | 1.41 | 0.04 | 0.71 | 1.24
Results from laboratory 6 were removed due to their 0.1 wt.% reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach (05%)
| Analytical results for Co in OREAS 554 (Certified Value 0.555 wt.%) | [20 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 07 | 10 | 12 | 18 | 19 | 20 | 23 | 24 | [10 empty] | wt.%
|  |  | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | [10 empty] | 3
3
| 1 | 1 | 0.5684 | 0.34 | 0.588 | 0.494 | 0.5441 | 0.546 | 0.566 | 0.595 | > 0.5 | [10 empty] | 1
| 1 | 2 | 0.5668 | 0.35 | 0.598 | 0.475 | 0.5141 | 0.5428 | 0.573 | 0.61 | > 0.5 | [10 empty] | 0
| 1 | 3 | 0.5858 | 0.34 | 0.584 | 0.481 | 0.5059 | 0.5247 | 0.562 | 0.591 | > 0.5 | [10 empty] | 16
| 1 | 4 | 0.6034 | 0.34 | 0.599 | 0.488 | 0.5307 | 0.5402 | 0.612 | 0.582 | > 0.5 | [10 empty] | 0.55549
| 1 | 5 | 0.5743 | 0.34 | 0.591 | 0.468 | 0.5186 | 0.5061 | 0.569 | 0.602 | > 0.5 | [10 empty] | 28
| 1 | 6 | 0.5865 | 0.34 | 0.587 | 0.484 | 0.5346 | 0.5324 | 0.606 | 0.6 | > 0.5
| Mean |  | 0.58087 | 0.34167 | 0.59117 | 0.48167 | 0.52467 | 0.53203 | 0.58133 | 0.59667
| Median |  | 0.58005 | 0.34 | 0.5895 | 0.4825 | 0.52465 | 0.5363 | 0.571 | 0.5975
| Std Dev. |  | 0.013858 | 0.0040825 | 0.0061128 | 0.0092664 | 0.014236 | 0.01486 | 0.021814 | 0.0096678
| Rel.Std.Dev. |  | 2.386% | 1.195% | 1.034% | 1.924% | 2.713% | 2.793% | 3.752% | 1.62%
| PDM3 |  | 4.569% | -38.49% | 6.423% | -13.29% | -5.548% | -4.222% | 4.653% | 7.414%
| Z-Score (Absolute) |  | 0.5 | 4.36 | 0.7 | 1.52 | 0.64 | 0.5 | 0.5 | 0.82 | NA
| Analytical results for Cu in OREAS 554 (Indicative Value 0.25 wt.%) | [20 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 05 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | wt.%
|  |  | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*OES | AL-5S*OES | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES/AAS | AL-5S*OES/AAS | 3
3
| 1 | 1 | 0.3071 | 0.2424 | 0.46 | 0.53 | 0.256 | 0.719 | 0.455 | 0.198 | 0.254 | 0.246 | 0.244 | 0.2437 | 1.195 | 0.696 | 0.28 | 0.833 | 0.232 | 0.249 | 0.253 | 1
| 1 | 2 | 0.2964 | 0.247 | 0.46 | 0.51915 | 0.253 | 0.721 | 0.454 | 0.203 | 0.255 | 0.25 | 0.244 | 0.2386 | 1.223 | 0.676 | 0.27 | 0.862 | 0.236 | 0.251 | 0.26 | 0
| 1 | 3 | 0.2889 | 0.2461 | 0.455 | 0.53571 | 0.259 | 0.714 | 0.468 | 0.198 | 0.255 | 0.246 | 0.253 | 0.2384 | 1.234 | 0.721 | 0.28 | 0.872 | 0.231 | 0.254 | 0.254 | 16
| 1 | 4 | 0.3157 | 0.2481 | 0.47 | 0.535 | 0.262 | 0.72 | 0.447 | 0.192 | 0.26 | 0.25 | 0.236 | 0.241 | 1.256 | 0.699 | 0.27 | 0.803 | 0.23 | 0.243 | 0.253 | 0.25038
| 1 | 5 | 0.3196 | 0.2499 | 0.49 | 0.53 | 0.26 | 0.727 | 0.434 | 0.193 | 0.262 | 0.249 | 0.245 | 0.2335 | 1.193 | 0.729 | 0.27 | 0.863 | 0.229 | 0.242 | 0.254 | 15
| 1 | 6 | 0.3028 | 0.2413 | 0.46 | 0.42984 | 0.266 | 0.721 | 0.45 | 0.201 | 0.249 | 0.248 | 0.241 | 0.2367 | 1.181 | 0.696 | 0.27 | 0.87 | 0.23 | 0.26 | 0.261
| Mean |  | 0.30508 | 0.2458 | 0.46583 | 0.51328 | 0.25933 | 0.72033 | 0.45133 | 0.1975 | 0.25583 | 0.24817 | 0.24383 | 0.23865 | 1.2137 | 0.70283 | 0.27333 | 0.8505 | 0.23133 | 0.24983 | 0.25583
| Median |  | 0.30495 | 0.24655 | 0.46 | 0.53 | 0.2595 | 0.7205 | 0.452 | 0.198 | 0.255 | 0.2485 | 0.244 | 0.2385 | 1.209 | 0.6975 | 0.27 | 0.8625 | 0.2305 | 0.25 | 0.254
| Std Dev. |  | 0.011576 | 0.0033299 | 0.012813 | 0.041305 | 0.0045461 | 0.0041793 | 0.01113 | 0.0043243 | 0.0046224 | 0.0018348 | 0.0055648 | 0.0035047 | 0.028759 | 0.019198 | 0.005164 | 0.027179 | 0.0025033 | 0.0067946 | 0.003656
| Rel.Std.Dev. |  | 3.794% | 1.355% | 2.751% | 8.047% | 1.753% | 0.5802% | 2.466% | 2.19% | 1.807% | 0.7394% | 2.282% | 1.469% | 2.37% | 2.732% | 1.889% | 3.196% | 1.082% | 2.72% | 1.429%
| PDM3 |  | 21.85% | -1.828% | 86.05% | 105% | 3.577% | 187.7% | 80.26% | -21.12% | 2.179% | -0.8831% | -2.614% | -4.684% | 384.7% | 180.7% | 9.168% | 239.7% | -7.606% | -0.2174% | 2.179%
| Z-Score (Absolute) |  | 1.1 | 0.33 | 4.97 | 6.12 | 0 | 11.1 | 4.62 | 1.49 | 0.08 | 0.27 | 0.37 | 0.5 | 22.98 | 10.68 | 0.34 | 14.24 | 0.67 | 0.23 | 0.08
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach (10%)
| Analytical results for Co in OREAS 554 (Certified Value 0.546 wt.%) | [13 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 07 | 10 | 11 | 12 | 18 | 19 | 20 | 23 | 24 | 26 | 27 | wt.%
|  |  | AL-10S*AAS | AL-10S*AAS | AL-10S*AAS | AL-10S*AAS | AL-10S*AAS | AL-10S*OES | AL-10S*OES | AL-10S*OES | AL-10S*AAS | AL-10S*AAS | AL-10S*OES | AL-10S*AAS | 3
3
| 1 | 1 | 0.563 | 0.34 | 0.587 | 0.53955 | 0.502 | 0.5141 | 0.5803 | 0.576 | 0.58 | 0.497 | 0.482 | 0.50628 | 1
| 1 | 2 | 0.5621 | 0.34 | 0.608 | 0.5269 | 0.504 | 0.5228 | 0.5837 | 0.609 | 0.58 | 0.487 | 0.479 | 0.49602 | 0
| 1 | 3 | 0.5585 | 0.34 | 0.596 | 0.52995 | 0.494 | 0.5258 | 0.5908 | 0.603 | 0.58 | > 0.5 | 0.5 | 0.47544 | 16
| 1 | 4 | 0.5695 | 0.34 | 0.607 | 0.53125 | 0.512 | 0.5253 | 0.603 | 0.583 | 0.579 | > 0.5 | 0.51 | 0.56215 | 0.54576
| 1 | 5 | 0.5583 | 0.33 | 0.607 | 0.50935 | 0.512 | 0.529 | 0.5978 | 0.571 | 0.58 | > 0.5 | 0.519 | 0.62457 | 30
| 1 | 6 | 0.5705 | 0.34 | 0.588 | 0.5181 | 0.497 | 1.001 | 0.5753 | 0.584 | 0.579 | > 0.5 | 0.494 | 0.59369
| Mean |  | 0.56365 | 0.33833 | 0.59883 | 0.52585 | 0.5035 | 0.603 | 0.58848 | 0.58767 | 0.57967 | 0.492 | 0.49733 | 0.54303
| Median |  | 0.56255 | 0.34 | 0.6015 | 0.52843 | 0.503 | 0.52555 | 0.58725 | 0.5835 | 0.58 | 0.492 | 0.497 | 0.53421
| Std Dev. |  | 0.0052748 | 0.0040825 | 0.0098268 | 0.01065 | 0.0074766 | 0.19504 | 0.010645 | 0.015095 | 0.0005164 | 0.0070711 | 0.015616 | 0.059514
| Rel.Std.Dev. |  | 0.9358% | 1.207% | 1.641% | 2.025% | 1.485% | 32.35% | 1.809% | 2.569% | 0.08909% | 1.437% | 3.14% | 10.96%
| PDM3 |  | 3.277% | -38.01% | 9.724% | -3.649% | -7.744% | 10.49% | 7.827% | 7.678% | 6.212% | -9.851% | -8.874% | -0.5017%
| Z-Score (Absolute) |  | 0.17 | 3.6 | 0.76 | 0.46 | 0.83 | 0.83 | 0.59 | 0.57 | 0.44 | 1.03 | 0.94 | 0.17
| Analytical results for Cu in OREAS 554 (Indicative Value 0.517 wt.%) | [13 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 02 | 07 | 10 | 11 | 12 | 18 | 19 | 20 | 23 | 24 | 26 | 27 | wt.%
|  |  | AL-10S*AAS | AL-10S*AAS | AL-10S*AAS | AL-10S*AAS | AL-10S*AAS | AL-10S*OES | AL-10S*OES | AL-10S*OES | AL-10S*AAS | AL-10S*AAS | AL-10S*OES | AL-10S*AAS | 3
3
| 1 | 1 | 0.6068 | 0.56 | 0.716 | 0.60495 | 0.558 | 0.2473 | 0.9727 | 0.739 | 0.72 | 0.278 | 0.192 | 0.18432 | 1
| 1 | 2 | 0.5799 | 0.55 | 0.722 | 0.5867 | 0.581 | 0.2485 | 0.9699 | 0.714 | 0.73 | 0.275 | 0.191 | 0.16833 | 0
| 1 | 3 | 0.3362 | 0.54 | 0.721 | 0.5915 | 0.547 | 0.2534 | 0.9964 | 0.731 | 0.72 | 0.269 | 0.192 | 0.17283 | 16
| 1 | 4 | 0.3496 | 0.53 | 0.72 | 0.58515 | 0.567 | 0.2545 | 0.9938 | 0.725 | 0.719 | 0.268 | 0.202 | 0.208 | 0.51706
| 1 | 5 | 0.3301 | 0.56 | 0.72 | 0.58435 | 0.57 | 0.256 | 0.998 | 0.74 | 0.72 | 0.273 | 0.208 | 0.23373 | 17
| 1 | 6 | 0.3202 | 0.57 | 0.721 | 0.5962 | 0.544 | 0.8073 | 0.9826 | 0.742 | 0.72 | 0.28 | 0.201 | 0.2183
| Mean |  | 0.42047 | 0.55167 | 0.72 | 0.59147 | 0.56117 | 0.3445 | 0.98557 | 0.73183 | 0.7215 | 0.27383 | 0.19767 | 0.19759
| Median |  | 0.3429 | 0.555 | 0.7205 | 0.5891 | 0.5625 | 0.25395 | 0.9882 | 0.735 | 0.72 | 0.274 | 0.1965 | 0.19616
| Std Dev. |  | 0.13452 | 0.01472 | 0.0020976 | 0.0079684 | 0.014219 | 0.22675 | 0.012325 | 0.010834 | 0.0041833 | 0.0047924 | 0.0070048 | 0.026414
| Rel.Std.Dev. |  | 31.99% | 2.668% | 0.2913% | 1.347% | 2.534% | 65.82% | 1.251% | 1.48% | 0.5798% | 1.75% | 3.544% | 13.37%
| PDM3 |  | -18.68% | 6.693% | 39.25% | 14.39% | 8.531% | -33.37% | 90.61% | 41.54% | 39.54% | -47.04% | -61.77% | -61.79%
| Z-Score (Absolute) |  | 0.54 | 0.02 | 0.65 | 0.14 | 0.02 | 0.84 | 1.7 | 0.69 | 0.65 | 1.12 | 1.42 | 1.42
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 554 (Certified Value < 1 ppm) | [22 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 18 | 20 | 21 | 29 | 30 | [12 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*OES | [12 empty] | 2
2
| 1 | 1 | < 1 | < 5 | < 5 | N.A. | NA | < 1 | < 1 | < 1 | 0.2 | [12 empty] | 1
| 1 | 2 | < 1 | < 5 | < 5 | N.A. | NA | < 1 | < 1 | < 1 | < 0.1 | [12 empty] | 0
| 1 | 3 | < 1 | < 5 | < 5 | N.A. | NA | < 1 | < 1 | < 1 | 0.2 | [12 empty] | 16
| 1 | 4 | < 1 | < 5 | < 5 | N.A. | NA | < 1 | < 1 | < 1 | < 0.1 | [12 empty] | < 1
| 1 | 5 | < 1 | < 5 | < 5 | N.A. | NA | < 1 | < 1 | < 1 | < 0.1 | [12 empty] | 32
| 1 | 6 | < 1 | < 5 | < 5 | N.A. | NA | < 1 | < 1 | < 1 | < 0.1
| Mean | [9 empty] | 0.2
| Median | [9 empty] | 0.2
| Std Dev. | [9 empty] | 0
| Rel.Std.Dev. | [9 empty] | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | NA | NA
| Analytical results for Al in OREAS 554 (Certified Value 4.32 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  |  | 2
3
| 1 | 1 | 4.19 | 4.24 | 4.22 | 4.45 | 4.68 | 4.44 | 4.94 | 4.282 | 4.308 | 4.271 | 4.43 | 4.3388 | 4.3706 | 4.18 | 4.4255 | 4.92 | 4.28 | 4.07 | 4.53 |  |  | 1
| 1 | 2 | 4.24 | 4.22 | 4.3 | 4.42 | 4.62 | 4.46 | 4.98 | 4.292 | 4.361 | 4.25 | 4.329 | 4.2618 | 4.3248 | 4.4 | 4.5392 | 5.07 | 4.28 | 4.02 | 4.488 |  |  | 0
| 1 | 3 | 4.24 | 4.28 | 4.39 | 4.4 | 4.53 | 4.47 | 5.08 | 4.308 | 4.26 | 4.387 | 4.361 | 4.309 | 4.232 | 4.59 | 4.2682 | 4.91 | 4.29 | 3.91 | 4.456 |  |  | 16
| 1 | 4 | 4.11 | 4.18 | 4.32 | 4.45 | 4.58 | 4.43 | 5.1 | 4.319 | 4.239 | 4.382 | 4.382 | 4.3002 | 4.2598 | 4.59 | 4.2608 | 4.9 | 4.1 | 4.02 | 4.424 |  |  | 4.3246
| 1 | 5 | 4.16 | 4.24 | 4.35 | 4.43 | 4.3 | 4.41 | 4.98 | 4.292 | 4.382 | 4.329 | 4.35 | 4.2861 | 4.2055 | 4.48 | 4.5769 | 4.94 | 4.11 | 4.13 | 4.472 |  |  | 33
| 1 | 6 | 4.16 | 4.26 | 4.27 | 4.4 | 4.29 | 4.46 | 5.11 | 4.282 | 4.297 | 4.34 | 4.372 | 4.1816 | 4.1783 | 4.54 | 4.2414 | 4.99 | 4.23 | 4.06 | 4.499
| Mean |  | 4.1833 | 4.2367 | 4.3083 | 4.425 | 4.5 | 4.445 | 5.0317 | 4.2958 | 4.3078 | 4.3265 | 4.3707 | 4.2796 | 4.2618 | 4.4633 | 4.3853 | 4.955 | 4.215 | 4.035 | 4.4782
| Median |  | 4.175 | 4.24 | 4.31 | 4.425 | 4.555 | 4.45 | 5.03 | 4.292 | 4.3025 | 4.3345 | 4.3665 | 4.2931 | 4.2459 | 4.51 | 4.3468 | 4.93 | 4.255 | 4.04 | 4.48
| Std Dev. |  | 0.050859 | 0.034448 | 0.059805 | 0.022583 | 0.16625 | 0.022583 | 0.073326 | 0.014811 | 0.055625 | 0.056323 | 0.034372 | 0.054326 | 0.073257 | 0.15655 | 0.14962 | 0.06473 | 0.087807 | 0.073417 | 0.036554
| Rel.Std.Dev. |  | 1.216% | 0.8131% | 1.388% | 0.5104% | 3.695% | 0.5081% | 1.457% | 0.3448% | 1.291% | 1.302% | 0.7864% | 1.269% | 1.719% | 3.507% | 3.412% | 1.306% | 2.083% | 1.819% | 0.8163%
| PDM3 |  | -3.266% | -2.032% | -0.3753% | 2.323% | 4.057% | 2.785% | 16.35% | -0.6643% | -0.3868% | 0.04483% | 1.066% | -1.04% | -1.451% | 3.209% | 1.405% | 14.58% | -2.533% | -6.696% | 3.552%
| Z-Score (Absolute) |  | 0.98 | 0.62 | 0.12 | 0.67 | 1.19 | 0.81 | 4.83 | 0.21 | 0.13 | 0 | 0.3 | 0.32 | 0.44 | 0.94 | 0.4 | 4.3 | 0.76 | 2 | 1.04
| Analytical results for As in OREAS 554 (Certified Value 66 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  |  | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES |  |  | 0
1
| 1 | 1 | 65 | 70 | 65 | 70 | < 100 | 71 | 72.6 | 100 | 100 | 100 | 100 | NA | 26.4426 | 66 | 54 | 80 | < 100 | 59 | 14.1 |  |  | 1
| 1 | 2 | 66 | 60 | 60 | 60 | < 100 | 74 | 70.2 | 100 | 200 | 100 | 200 | NA | 26.8404 | 65 | 58 | 70 | < 100 | 58 | 14.6 |  |  | 0
| 1 | 3 | 67 | 70 | 70 | 70 | < 100 | 73 | 71 | 100 | 200 | 100 | 100 | NA | 22.6182 | 66 | 54 | 80 | < 100 | 64 | 18.8 |  |  | 16
| 1 | 4 | 65 | 60 | 70 | 70 | < 100 | 70 | 71.2 | < 100 | 200 | 100 | 100 | NA | 18.8273 | 64 | 58 | 80 | < 100 | 63 | 20.5 |  |  | 65.869
| 1 | 5 | 66 | 70 | 70 | 70 | < 100 | 69 | 72.2 | 100 | 100 | 100 | 200 | NA | 22.7073 | 65 | 68 | 70 | < 100 | 64 | 19.8 |  |  | 34
| 1 | 6 | 66 | 70 | 70 | 70 | < 100 | 71 | 70.7 | < 100 | 200 | 100 | 200 | NA | 21.4652 | 66 | 59 | 70 | < 100 | 62 | 18.9
| Mean |  | 65.8333 | 66.6667 | 67.5 | 68.3333 |  | 71.3333 | 71.3167 | 100 | 166.6667 | 100 | 150 |  | 23.1502 | 65.3333 | 58.5 | 75 |  | 61.6667 | 17.7833
| Median |  | 66 | 70 | 70 | 70 |  | 71 | 71.1 | 100 | 200 | 100 | 150 |  | 22.6628 | 65.5 | 58 | 75 |  | 62.5 | 18.85
| Std Dev. |  | 0.75277 | 5.164 | 4.1833 | 4.0825 |  | 1.8619 | 0.91305 | 0 | 51.6398 | 0 | 54.7723 |  | 3.0484 | 0.8165 | 5.1284 | 5.4772 |  | 2.582 | 2.736
| Rel.Std.Dev. |  | 1.143% | 7.746% | 6.197% | 5.974% |  | 2.61% | 1.28% | 0% | 30.98% | 0% | 36.51% |  | 13.17% | 1.25% | 8.766% | 7.303% |  | 4.187% | 15.39%
| PDM3 |  | -0.05422% | 1.211% | 2.476% | 3.741% |  | 8.296% | 8.27% | 51.82% | 153% | 51.82% | 127.7% |  | -64.85% | -0.8133% | -11.19% | 13.86% |  | -6.38% | -73%
| Z-Score (Absolute) |  | 0.2 | NA | 0.34 | NA | 1.16 | 0.68 | 0.67 | 1.71 | NA | NA | NA | NA | 3.48 | 0.16 | 0.43 | NA | 1.16 | 0.16 | 3.94
Results from laboratories 3, 6, 14, 15, 16 & 23 were removed due to their 10 ppm or greater reading resolution.
| Analytical results for B in OREAS 554 (Certified Value 178 ppm) | [22 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [14 empty] | 1
| No | No | 02 | 03 | 06 | 10 | 20 | 21 | 23 | [14 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | [14 empty] | 0
0
| 1 | 1 | 157 | 160 | 150 | 170 | 194 | < 1000 | 190 | [14 empty] | 1
| 1 | 2 | 155 | 180 | 200 | 170 | 213 | < 1000 | 190 | [14 empty] | 0
| 1 | 3 | 154 | 180 | 150 | 160 | 209 | < 1000 | 190 | [14 empty] | 16
| 1 | 4 | 157 | 180 | 200 | 170 | 201 | < 1000 | 190 | [14 empty] | 178.3333
| 1 | 5 | 154 | 180 | 200 | 160 | 207 | < 1000 | 190 | [14 empty] | 35
| 1 | 6 | 162 | 180 | 150 | 170 | 207 | < 1000 | 190
| Mean |  | 156.5 | 176.6667 | 175 | 166.6667 | 205.1667 |  | 190
| Median |  | 156 | 180 | 175 | 170 | 207 |  | 190
| Std Dev. |  | 3.0166 | 8.165 | 27.3861 | 5.164 | 6.7057 |  | 0
| Rel.Std.Dev. |  | 1.928% | 4.622% | 15.65% | 3.098% | 3.268% |  | 0%
| PDM3 |  | -12.24% | -0.9346% | -1.869% | -6.542% | 15.05% |  | 6.542%
| Z-Score (Absolute) |  | 1.02 | 0 | 0.08 | 0.51 | 1.44 | 16.35 | 0.67
| Analytical results for Ba in OREAS 554 (Certified Value 1280 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 | [5 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | [5 empty] | 0
0
| 1 | 1 | 1331 | 1302 | 1260 | 1290 | 1300 | 1340 | 1370 | 1250 | 1136.5315 | 1230 | 1237 | 1340 | 804 | 1202 | 1125.4 | 1360 | [5 empty] | 1
| 1 | 2 | 1342 | 1294 | 1250 | 1270 | 1300 | 1410 | 1360 | 1234 | 1191.1608 | 1270 | 1255 | 1390 | 1216 | 1223 | 1230.5 | 1420 | [5 empty] | 0
| 1 | 3 | 1335 | 1344 | 1260 | 1280 | 1300 | 1420 | 1330 | 1254 | 1184.9383 | 1270 | 1230 | 1330 | 1226 | 1238 | 1248.5 | 1410 | [5 empty] | 16
| 1 | 4 | 1197 | 1267 | 1290 | 1280 | 1300 | 1360 | 1360 | 1238 | 1192.5328 | 1260 | 1237 | 1400 | 1172 | 1221 | 1189.4 | 1320 | [5 empty] | 1280.486
| 1 | 5 | 1342 | 1262 | 1240 | 1270 | 1300 | 1320 | 1360 | 1237 | 1201.1769 | 1260 | 1298 | 1340 | 1181 | 1211 | 1158 | 1380 | [5 empty] | 36
| 1 | 6 | 1338 | 1323 | 1290 | 1280 | 1300 | 1380 | 1330 | 1223 | 1194.8153 | 1260 | 1244 | 1340 | 1218 | 1201 | 1217.5 | 1410
| Mean |  | 1314.1667 | 1298.6667 | 1265 | 1278.3333 | 1300 | 1371.6667 | 1351.6667 | 1239.3333 | 1183.5259 | 1258.3333 | 1250.1667 | 1356.6667 | 1136.1667 | 1216 | 1194.8833 | 1383.3333
| Median |  | 1336.5 | 1298 | 1260 | 1280 | 1300 | 1370 | 1360 | 1237.5 | 1191.8468 | 1260 | 1240.5 | 1340 | 1198.5 | 1216 | 1203.45 | 1395
| Std Dev. |  | 57.5549 | 31.7091 | 20.7364 | 7.5277 | 0 | 39.2003 | 17.224 | 11.2368 | 23.6173 | 14.7196 | 24.9112 | 30.1109 | 164.176 | 14.1704 | 46.633 | 38.2971
| Rel.Std.Dev. |  | 4.38% | 2.442% | 1.639% | 0.5889% | 0% | 2.858% | 1.274% | 0.9067% | 1.996% | 1.17% | 1.993% | 2.219% | 14.45% | 1.165% | 3.903% | 2.768%
| PDM3 |  | 2.63% | 1.42% | -1.209% | -0.1681% | 1.524% | 7.121% | 5.559% | -3.214% | -7.572% | -1.73% | -2.368% | 5.949% | -11.27% | -5.036% | -6.685% | 8.032%
| Z-Score (Absolute) |  | 0.58 | 0.37 | 0.09 | 0.09 | 0.39 | 1.37 | 1.1 | 0.44 | 1.21 | 0.18 | 0.3 | 1.17 | 1.86 | 0.76 | 1.05 | 1.53
| Analytical results for Be in OREAS 554 (Indicative Value 2.22 ppm) | [22 empty] | Indicative
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 29 | 30 | [8 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | [8 empty] | 2
2
| 1 | 1 | < 5 | 2 | 2 | 2 | 3 | < 3 | < 5 | NA | 2.4472 | < 5 | < 5 | < 5 | 1.9 | [8 empty] | 1
| 1 | 2 | < 5 | 2 | 2 | 2 | 3 | < 3 | < 5 | NA | 2.1656 | < 5 | < 5 | < 5 | 4.9 | [8 empty] | 0
| 1 | 3 | < 5 | 1 | 2 | 2 | 2 | < 3 | < 5 | NA | 2.2511 | < 5 | < 5 | < 5 | 3.3 | [8 empty] | 16
| 1 | 4 | < 5 | 2 | 2 | 2 | 2 | < 3 | < 5 | NA | 2.3147 | < 5 | < 5 | < 5 | 1.2 | [8 empty] | 2.2185
| 1 | 5 | < 5 | 2 | 2 | 2 | 3 | < 3 | < 5 | NA | 2.1935 | < 5 | < 5 | < 5 | 3 | [8 empty] | 19
| 1 | 6 | < 5 | 1 | 2 | 2 | 3 | < 3 | < 5 | NA | 2.2955 | < 5 | < 5 | < 5 | 4.1
| Mean |  |  | 1.6667 | 2 | 2 | 2.6667 |  |  |  | 2.2779 |  |  |  | 3.0667
| Median |  |  | 2 | 2 | 2 | 3 |  |  |  | 2.2733 |  |  |  | 3.15
| Std Dev. |  |  | 0.5164 | 0 | 0 | 0.5164 |  |  |  | 0.10076 |  |  |  | 1.3663
| Rel.Std.Dev. |  |  | 30.98% | 0% | 0% | 19.36% |  |  |  | 4.423% |  |  |  | 44.55%
| PDM3 |  |  | -24.88% | -9.851% | -9.851% | 20.2% |  |  |  | 2.677% |  |  |  | 38.23%
| Z-Score (Absolute) |  | 0 | 2.89 | 1.73 | 1.73 | 0.58 | 3.47 | 0 | NA | 0.77 | 0 | 0 | 0 | 1.97
| Analytical results for Bi in OREAS 554 (Certified Value 2.83 ppm) | [22 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 18 | 19 | 20 | 21 | 23 | 29 | 30 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | [8 empty] | 2
2
| 1 | 1 | 3.1 | 3 | 2 | 3 | 3 | 3 | NA | 1.4201 | 3.6 | 2.7 | 3.6 | 2.5 | 1.2 | [8 empty] | 1
| 1 | 2 | 3 | 2.8 | 2 | 3 | 3 | 3 | NA | 1.3559 | 3.9 | 2.5 | 3.8 | 2.5 | 0.4 | [8 empty] | 0
| 1 | 3 | 2.9 | 3 | 2 | 3 | 3 | 3 | NA | 1.2475 | 3.1 | 2.6 | 4 | 2.4 | 2.3 | [8 empty] | 16
| 1 | 4 | 2.8 | 3 | 2 | 3 | 3.5 | 3 | NA | 1.4071 | 2.9 | 2.6 | 3.8 | 2.6 | 2.4 | [8 empty] | 2.8264
| 1 | 5 | 2.8 | 3.1 | 2 | 3 | 3 | 3 | NA | 1.2851 | 3.2 | 3.3 | 3.6 | 2.5 | 2.3 | [8 empty] | 37
| 1 | 6 | 2.9 | 2.9 | 2 | 3 | 3 | 3 | NA | 1.1566 | 3.2 | 2.7 | 3.7 | 2.5 | 1.9
| Mean |  | 2.9167 | 2.9667 | 2 | 3 | 3.0833 | 3 |  | 1.3121 | 3.3167 | 2.7333 | 3.75 | 2.5 | 1.75
| Median |  | 2.9 | 3 | 2 | 3 | 3 | 3 |  | 1.3205 | 3.2 | 2.65 | 3.75 | 2.5 | 2.1
| Std Dev. |  | 0.1169 | 0.10328 | 0 | 0 | 0.20412 | 0 |  | 0.10166 | 0.3656 | 0.28752 | 0.15166 | 0.063246 | 0.79687
| Rel.Std.Dev. |  | 4.008% | 3.481% | 0% | 0% | 6.62% | 0% |  | 7.748% | 11.02% | 10.52% | 4.044% | 2.53% | 45.54%
| PDM3 |  | 3.195% | 4.964% | -29.24% | 6.143% | 9.092% | 6.143% |  | -53.58% | 17.35% | -3.292% | 32.68% | -11.55% | -38.08%
| Z-Score (Absolute) |  | 0.06 | 0.06 | 2.18 | 0.13 | 0.33 | 0.13 | NA | 3.77 | 0.87 | 0.48 | 1.87 | 1.02 | 2.76
| Analytical results for Ca in OREAS 554 (Certified Value 3.96 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | BF*OES/MS |  | 2
3
| 1 | 1 | 4 | 3.9 | 3.9 | 3.8 | 4.1 | 4 | 3.8 | 3.74 | 4.002 | 3.902 | 4.052 | 4.331 | 4.3053 | 3.9644 | 3.91 | 3.7765 | 4.2 | 3.81 | 4.05 | 4.067 |  | 1
| 1 | 2 | 4 | 3.9 | 3.9 | 3.8 | 4 | 3.81 | 3.82 | 3.83 | 3.938 | 4.067 | 4.059 | 4.202 | 4.4546 | 3.9194 | 4.09 | 3.8722 | 4.2 | 3.81 | 3.81 | 4.074 |  | 0
| 1 | 3 | 4.1 | 4 | 4 | 3.8 | 4.1 | 3.59 | 3.84 | 3.71 | 4.038 | 3.881 | 4.081 | 4.217 | 4.3858 | 3.9115 | 4.11 | 3.75 | 4.2 | 3.86 | 3.95 | 4.031 |  | 16
| 1 | 4 | 4 | 3.8 | 3.9 | 3.9 | 4 | 3.71 | 3.81 | 3.75 | 3.981 | 3.859 | 4.038 | 4.231 | 4.5941 | 3.9344 | 4.09 | 3.71 | 4.2 | 3.86 | 3.9 | 4.045 |  | 3.9556
| 1 | 5 | 4 | 4 | 4.1 | 3.8 | 4.1 | 3.56 | 3.78 | 3.69 | 3.959 | 4.088 | 4.038 | 4.231 | 4.519 | 3.9301 | 4.09 | 3.5009 | 4.3 | 3.73 | 3.99 | 4.052 |  | 38
| 1 | 6 | 4 | 4 | 4 | 3.9 | 4 | 3.63 | 3.82 | 3.76 | 3.988 | 4.009 | 4.102 | 4.231 | 4.4118 | 3.9137 | 4.2 | 3.6953 | 4.3 | 3.79 | 4.07 | 4.088
| Mean |  | 4.0167 | 3.9333 | 3.9667 | 3.8333 | 4.05 | 3.7167 | 3.8117 | 3.7467 | 3.9843 | 3.9677 | 4.0617 | 4.2405 | 4.4451 | 3.9289 | 4.0817 | 3.7175 | 4.2333 | 3.81 | 3.9617 | 4.0595
| Median |  | 4 | 3.95 | 3.95 | 3.8 | 4.05 | 3.67 | 3.815 | 3.745 | 3.9845 | 3.9555 | 4.0555 | 4.231 | 4.4332 | 3.9247 | 4.09 | 3.73 | 4.2 | 3.81 | 3.97 | 4.0595
| Std Dev. |  | 0.040825 | 0.08165 | 0.08165 | 0.05164 | 0.054772 | 0.16561 | 0.020412 | 0.048442 | 0.034656 | 0.099687 | 0.025382 | 0.045816 | 0.1019 | 0.019583 | 0.094322 | 0.12323 | 0.05164 | 0.04858 | 0.097245 | 0.020773
| Rel.Std.Dev. |  | 1.016% | 2.076% | 2.058% | 1.347% | 1.352% | 4.456% | 0.5355% | 1.293% | 0.8698% | 2.512% | 0.6249% | 1.08% | 2.292% | 0.4984% | 2.311% | 3.315% | 1.22% | 1.275% | 2.455% | 0.5117%
| PDM3 |  | 1.544% | -0.5629% | 0.2798% | -3.091% | 2.387% | -6.04% | -3.639% | -5.282% | 0.7264% | 0.3051% | 2.681% | 7.202% | 12.37% | -0.6747% | 3.187% | -6.02% | 7.021% | -3.681% | 0.1534% | 2.627%
| Z-Score (Absolute) |  | 0.32 | 0.22 | 0 | 0.86 | 0.53 | 1.62 | 1 | 1.42 | 0.11 | 0 | 0.61 | 1.76 | 3.08 | 0.25 | 0.74 | 1.61 | 1.72 | 1.01 | 0.04 | 0.6
| Analytical results for Cd in OREAS 554 (Indicative Value 0.23 ppm) | [22 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 24 | 29 | 30 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | [8 empty] | 2
2
| 1 | 1 | < 0.2 | < 10 | < 10 | < 1 | < 2 | 0.3 | NA | 1.2671 | < 0.5 | < 10 | < 100 | 0.2 | 2.2 | [8 empty] | 1
| 1 | 2 | 0.2 | < 10 | < 10 | < 1 | < 2 | 0.3 | NA | 1.1497 | < 0.5 | < 10 | < 100 | 0.3 | 2.3 | [8 empty] | 0
| 1 | 3 | < 0.2 | < 10 | < 10 | < 1 | < 2 | 0.3 | NA | 1.8502 | < 0.5 | < 10 | < 100 | 0.2 | 1.5 | [8 empty] | 16
| 1 | 4 | 0.2 | < 10 | < 10 | < 1 | < 2 | 0.4 | NA | 1.5135 | < 0.5 | < 10 | < 100 | < 0.2 | 1.3 | [8 empty] | 0.22778
| 1 | 5 | < 0.2 | < 10 | < 10 | < 1 | < 2 | 0.3 | NA | 1.6971 | < 0.5 | < 10 | < 100 | 0.2 | 13.6 | [8 empty] | 20
| 1 | 6 | < 0.2 | < 10 | < 10 | < 1 | < 2 | 0.4 | NA | 2.1709 | < 0.5 | < 10 | < 100 | 0.3 | 0.7
| Mean |  | 0.2 |  |  |  |  | 0.33333 |  | 1.6081 |  |  |  | 0.24 | 3.6
| Median |  | 0.2 |  |  |  |  | 0.3 |  | 1.6053 |  |  |  | 0.2 | 1.85
| Std Dev. |  | 0 |  |  |  |  | 0.05164 |  | 0.37912 |  |  |  | 0.054772 | 4.9348
| Rel.Std.Dev. |  | 0% |  |  |  |  | 15.49% |  | 23.58% |  |  |  | 22.82% | 137.1%
| PDM3 |  | -12.2% |  |  |  |  | 46.34% |  | 606% |  |  |  | 5.366% | 1480%
| Z-Score (Absolute) |  | 0.7 | 2.21 | 2.21 | 0.48 | 0.18 | 0.58 | NA | 0.18 | 0.63 | 2.21 | 29.08 | 0.65 | 1.37
| Analytical results for Ce in OREAS 554 (Certified Value 68 ppm) | [22 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 10 | 11 | 18 | 21 | 28 | 29 | 30 | [10 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | PF*MS | PF*OES | [10 empty] | 0
1
| 1 | 1 | 72.2 | 70 | 65.5 | 68 | 70 | 69.9 | NA | 60.5 | 62.7 | 51.6 | 68 | [10 empty] | 1
| 1 | 2 | 69.2 | 70 | 66.5 | 67 | 68.4 | 69.4 | NA | 60.2 | 63.6 | 52.6 | 67 | [10 empty] | 0
| 1 | 3 | 69.3 | 68 | 66 | 68 | 70.2 | 68.7 | NA | 59.5 | 64.9 | 52.8 | 67 | [10 empty] | 16
| 1 | 4 | 68.3 | 69.5 | 66 | 68 | 75.3 | 68.6 | NA | 60.3 | 63.6 | 53.4 | 73 | [10 empty] | 67.7987
| 1 | 5 | 71 | 68.5 | 65 | 69 | 68.5 | 68.6 | NA | 62.1 | 64.2 | 54 | 70 | [10 empty] | 39
| 1 | 6 | 70.4 | 69 | 65.5 | 66 | 69.6 | 68.1 | NA | 58.6 | 62.1 | 51.5 | 68
| Mean |  | 70.0667 | 69.1667 | 65.75 | 67.6667 | 70.3333 | 68.8833 |  | 60.2 | 63.5167 | 52.65 | 68.8333
| Median |  | 69.85 | 69.25 | 65.75 | 68 | 69.8 | 68.65 |  | 60.25 | 63.6 | 52.7 | 68
| Std Dev. |  | 1.4137 | 0.8165 | 0.5244 | 1.0328 | 2.5469 | 0.64936 |  | 1.1628 | 1.0068 | 0.98336 | 2.3166
| Rel.Std.Dev. |  | 2.018% | 1.18% | 0.7976% | 1.526% | 3.621% | 0.9427% |  | 1.931% | 1.585% | 1.868% | 3.366%
| PDM3 |  | 3.345% | 2.018% | -3.022% | -0.1948% | 3.738% | 1.6% |  | -11.21% | -6.316% | -22.34% | 1.526%
| Z-Score (Absolute) |  | 0.63 | 0.32 | 0.86 | 0.2 | 0.72 | 0.22 | NA | 2.78 | 1.64 | 5.39 | 0.2
| Analytical results for Co in OREAS 554 (Certified Value 0.809 wt.%) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 | wt.%
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | 3
3
| 1 | 1 | 0.7602 | 0.8004 | 0.778 | 0.816 | 0.795 | 0.86 | > 0.5 | 0.856 | 0.802 | 0.799 | 0.823 | 0.828 | 0.85211 | 0.853 | 0.813 | 0.7947 | 0.817 | > 0.5 | 0.8131 | 0.75938 | 0.806 | 1
| 1 | 2 | 0.7214 | 0.8011 | 0.788 | 0.816 | 0.785 | 0.815 | > 0.5 | 0.863 | 0.819 | 0.807 | 0.823 | 0.816 | 0.84334 | 0.843 | 0.827 | 0.8104 | 0.811 | > 0.5 | 0.8097 | 0.75933 | 0.858 | 0
| 1 | 3 | 0.8226 | 0.8024 | 0.81 | 0.818 | 0.79 | 0.817 | > 0.5 | 0.851 | 0.817 | 0.811 | 0.827 | 0.806 | 0.84124 | 0.839 | 0.835 | 0.7887 | 0.809 | > 0.5 | 0.7832 | 0.76685 | 0.842 | 16
| 1 | 4 | 0.8021 | 0.7799 | 0.816 | 0.82 | 0.8 | 0.819 | > 0.5 | 0.866 | 0.805 | 0.791 | 0.816 | 0.827 | 0.83131 | 0.842 | 0.814 | 0.8815 | 0.808 | > 0.5 | 0.7896 | 0.76072 | 0.805 | 0.80949
| 1 | 5 | 0.7481 | 0.8079 | 0.83 | 0.816 | 0.8 | 0.79 | > 0.5 | 0.865 | 0.8 | 0.823 | 0.831 | 0.845 | 0.85011 | 0.865 | 0.823 | 0.773 | 0.799 | > 0.5 | 0.8102 | 0.75984 | 0.824 | 40
| 1 | 6 | 0.7894 | 0.8121 | 0.808 | 0.81 | 0.8 | 0.771 | > 0.5 | 0.861 | 0.813 | 0.806 | 0.847 | 0.807 | 0.83434 | 0.847 | 0.824 | 0.7813 | 0.797 | > 0.5 | 0.7778 | 0.76721 | 0.809
| Mean |  | 0.77397 | 0.80063 | 0.805 | 0.816 | 0.795 | 0.812 |  | 0.86033 | 0.80933 | 0.80617 | 0.82783 | 0.8215 | 0.84208 | 0.84817 | 0.82267 | 0.80493 | 0.80683 |  | 0.79727 | 0.76222 | 0.824
| Median |  | 0.7748 | 0.80175 | 0.809 | 0.816 | 0.7975 | 0.816 |  | 0.862 | 0.809 | 0.8065 | 0.825 | 0.8215 | 0.84229 | 0.845 | 0.8235 | 0.7917 | 0.8085 |  | 0.79965 | 0.76028 | 0.8165
| Std Dev. |  | 0.037461 | 0.011114 | 0.018963 | 0.0033466 | 0.0063246 | 0.030173 |  | 0.005785 | 0.0080664 | 0.010852 | 0.010629 | 0.014869 | 0.0082858 | 0.0095586 | 0.0082624 | 0.039595 | 0.0075476 |  | 0.015544 | 0.0037595 | 0.021863
| Rel.Std.Dev. |  | 4.84% | 1.388% | 2.356% | 0.4101% | 0.7955% | 3.716% |  | 0.6724% | 0.9967% | 1.346% | 1.284% | 1.81% | 0.984% | 1.127% | 1.004% | 4.919% | 0.9355% |  | 1.95% | 0.4932% | 2.653%
| PDM3 |  | -4.388% | -1.094% | -0.5545% | 0.8044% | -1.79% | 0.3102% |  | 6.281% | -0.01919% | -0.4104% | 2.266% | 1.484% | 4.026% | 4.778% | 1.628% | -0.5627% | -0.328% |  | -1.51% | -5.839% | 1.793%
| Z-Score (Absolute) |  | 1.96 | 0.48 | 0.24 | 0.37 | 0.79 | 0.15 | NA | 2.83 | 0 | 0.18 | 1.03 | 0.67 | 1.81 | 2.15 | 0.74 | 0.24 | 0.14 | NA | 0.67 | 2.61 | 0.81
| Analytical results for Cr in OREAS 554 (Certified Value 69 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  |  | 0
0
| 1 | 1 | 56 | 79 | 50 | 60 | 50 | 100 | 100 | 100 | 100 | 70 | 140 | NA | 73.9785 | 70 | 54 | 90 | < 100 | 53 | 61 |  |  | 1
| 1 | 2 | 53 | 83 | 50 | 70 | 100 | 300 | 140 | 100 | 100 | 70 | 70 | NA | 70.9946 | 70 | 87 | 80 | < 100 | 51 | 72 |  |  | 0
| 1 | 3 | 55 | 77 | < 50 | 70 | 100 | 200 | 110 | 100 | 200 | 70 | 70 | NA | 76.9404 | 70 | 54 | 80 | < 100 | 52 | 78 |  |  | 16
| 1 | 4 | 54 | 74 | < 50 | 70 | 100 | 100 | 110 | 100 | 200 | 70 | 70 | NA | 78.1652 | 70 | 56 | 90 | < 100 | 49 | 61 |  |  | 68.6766
| 1 | 5 | 53 | 78 | 50 | 70 | 100 | < 100 | 100 | 100 | 100 | 70 | 70 | NA | 76.1637 | 70 | 63 | 90 | < 100 | 51 | 65 |  |  | 41
| 1 | 6 | 54 | 83 | 50 | 60 | 100 | < 100 | 120 | 100 | 100 | 70 | 70 | NA | 80.6607 | 70 | 57 | 80 | < 100 | 52 | 64
| Mean |  | 54.1667 | 79 | 50 | 66.6667 | 91.6667 | 175 | 113.3333 | 100 | 133.3333 | 70 | 81.6667 |  | 76.1505 | 70 | 61.8333 | 85 |  | 51.3333 | 66.8333
| Median |  | 54 | 78.5 | 50 | 70 | 100 | 150 | 110 | 100 | 100 | 70 | 70 |  | 76.5521 | 70 | 56.5 | 85 |  | 51.5 | 64.5
| Std Dev. |  | 1.169 | 3.5214 | 0 | 5.164 | 20.4124 | 95.7427 | 15.0555 | 0 | 51.6398 | 0 | 28.5774 |  | 3.3564 | 0 | 12.7658 | 5.4772 |  | 1.3663 | 6.7946
| Rel.Std.Dev. |  | 2.158% | 4.457% | 0% | 7.746% | 22.27% | 54.71% | 13.28% | 0% | 38.73% | 0% | 34.99% |  | 4.408% | 0% | 20.65% | 6.444% |  | 2.662% | 10.17%
| PDM3 |  | -21.13% | 15.03% | -27.2% | -2.927% | 33.48% | 154.8% | 65.02% | 45.61% | 94.15% | 1.927% | 18.91% |  | 10.88% | 1.927% | -9.965% | 23.77% |  | -25.25% | -2.684%
| Z-Score (Absolute) |  | 0.84 | 0.63 | 1.58 | 0.1 | 1.37 | NA | 2.66 | NA | NA | 0 | 0.67 | NA | 0.46 | 0.09 | 0.39 | 0.98 | 1.09 | 1.01 | 0.09
Results from laboratories 9, 13 & 14 were removed due to their 100 ppm reading resolution.
| Analytical results for Cs in OREAS 554 (Certified Value 2.55 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [7 empty] | 2
3
| 1 | 1 | 2.3 | < 10 | 3 | 5 | 3.2 | 2.5 | NA | 2.4571 | 2.8 | 2.6 | 2.7 | 2.2 | 2.5 | 2.5 | [7 empty] | 1
| 1 | 2 | 2.5 | < 10 | 2 | 4 | 3.1 | 2.6 | NA | 2.2567 | 2.9 | 2.5 | 2.5 | 2.3 | 2.6 | 2.3 | [7 empty] | 0
| 1 | 3 | 2.5 | < 10 | 3 | 3 | 2.6 | 2.5 | NA | 2.1862 | 2.5 | 2.5 | 2.5 | 2.3 | 2.6 | 2.5 | [7 empty] | 16
| 1 | 4 | 2.4 | < 10 | 2 | 3 | 2.2 | 2.5 | NA | 2.1977 | 2.4 | 2.6 | 2.4 | 2.3 | 2.8 | 2.4 | [7 empty] | 2.5548
| 1 | 5 | 2.4 | < 10 | 3 | 4 | 4.3 | 2.5 | NA | 2.1905 | 2.5 | 2.6 | 2.5 | 2.7 | 2.7 | 2.6 | [7 empty] | 42
| 1 | 6 | 2.4 | < 10 | 2 | 4 | 2.6 | 2.6 | NA | 2.1715 | 2.6 | 2.5 | 2.7 | 2.2 | 2.8 | 2.5
| Mean |  | 01: 2.4167; 03: 2.5; 04: 3.8333; 06: 3; 09: 2.5333; 11: 2.2433; 13: 2.6167; 14: 2.55; 15: 2.55; 16: 2.3333; 18: 2.6667; 19: 2.4667
| Median |  | 01: 2.4; 03: 2.5; 04: 4; 06: 2.85; 09: 2.5; 11: 2.1941; 13: 2.55; 14: 2.55; 15: 2.5; 16: 2.3; 18: 2.65; 19: 2.5
| Std Dev. |  | 01: 0.075277; 03: 0.54772; 04: 0.75277; 06: 0.73485; 09: 0.05164; 11: 0.10877; 13: 0.19408; 14: 0.054772; 15: 0.12247; 16: 0.18619; 18: 0.12111; 19: 0.10328
| Rel.Std.Dev. |  | 01: 3.115%; 03: 21.91%; 04: 19.64%; 06: 24.49%; 09: 2.038%; 11: 4.849%; 13: 7.417%; 14: 2.148%; 15: 4.803%; 16: 7.98%; 18: 4.541%; 19: 4.187%
| PDM3 |  | 01: -5.405%; 03: -2.144%; 04: 50.05%; 06: 17.43%; 09: -0.8388%; 11: -12.19%; 13: 2.423%; 14: -0.1864%; 15: -0.1864%; 16: -8.667%; 18: 4.38%; 19: -3.448%
| Z-Score (Absolute) |  | NA | 19.82 | 0.4 | NA | 3.64 | 0.13 | NA | 2.48 | 0.54 | 0 | 0 | 1.75 | 0.94 | 0.67
Results from laboratories 1 & 6 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 554 (Certified Value 1.57 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 1.65 | 1.5537 | 1.6 | 1.6 | 1.55 | 1.604 | > 1 | 1.72 | 1.525 | 1.525 | 1.57 | 1.62 | 1.6072 | 1.6033 | 1.56 | 1.6185 | 1.62 | 1.5435 | 1.5063 | 1.32 |  | 1
| 1 | 2 | 1.72 | 1.5206 | 1.62 | 1.62 | 1.57 | 1.518 | > 1 | 1.73 | 1.56 | 1.565 | 1.57 | 1.59 | 1.535 | 1.5948 | 1.6 | 1.585 | 1.67 | 1.5435 | 1.5058 | 1.27 |  | 0
| 1 | 3 | 1.7 | 1.5814 | 1.57 | 1.61 | 1.55 | 1.546 | > 1 | 1.71 | 1.59 | 1.51 | 1.6 | 1.6 | 1.605 | 1.5459 | 1.59 | 1.5448 | 1.65 | 1.5557 | 1.4925 | 1.32 |  | 16
| 1 | 4 | 1.61 | 1.5105 | 1.55 | 1.58 | 1.55 | 1.57 | > 1 | 1.74 | 1.575 | 1.51 | 1.57 | 1.6 | 1.5728 | 1.5459 | 1.6 | 1.5619 | 1.64 | 1.4906 | 1.5084 | 1.13 |  | 1.5693
| 1 | 5 | 1.72 | 1.5453 | 1.56 | 1.6 | 1.57 | 1.485 | > 1 | 1.71 | 1.595 | 1.6 | 1.56 | 1.6 | 1.5876 | 1.5586 | 1.6 | 1.6467 | 1.65 | 1.4914 | 1.4893 | 1.25 |  | 43
| 1 | 6 | 1.69 | 1.583 | 1.57 | 1.63 | 1.56 | 1.486 | > 1 | 1.71 | 1.53 | 1.545 | 1.61 | 1.59 | 1.5496 | 1.5437 | 1.64 | 1.5363 | 1.65 | 1.5183 | 1.5174 | 1.06
| Mean |  | 1.6817 | 1.5491 | 1.5783 | 1.6067 | 1.5583 | 1.5348 |  | 1.72 | 1.5625 | 1.5425 | 1.58 | 1.6 | 1.5762 | 1.5654 | 1.5983 | 1.5822 | 1.6467 | 1.5238 | 1.5033 | 1.225
| Median |  | 1.695 | 1.5495 | 1.57 | 1.605 | 1.555 | 1.532 |  | 1.715 | 1.5675 | 1.535 | 1.57 | 1.6 | 1.5802 | 1.5522 | 1.6 | 1.5734 | 1.65 | 1.5309 | 1.5061 | 1.26
| Std Dev. |  | 0.043551 | 0.030101 | 0.026394 | 0.017512 | 0.0098319 | 0.047533 |  | 0.012649 | 0.029791 | 0.03532 | 0.02 | 0.010954 | 0.029447 | 0.02676 | 0.025626 | 0.043386 | 0.01633 | 0.028199 | 0.010514 | 0.10672
| Rel.Std.Dev. |  | 2.59% | 1.943% | 1.672% | 1.09% | 0.6309% | 3.097% |  | 0.7354% | 1.907% | 2.29% | 1.266% | 0.6847% | 1.868% | 1.71% | 1.603% | 2.742% | 0.9917% | 1.851% | 0.6994% | 8.712%
| PDM3 |  | 7.163% | -1.286% | 0.5783% | 2.384% | -0.6962% | -2.194% |  | 9.606% | -0.4307% | -1.705% | 0.6845% | 1.959% | 0.4436% | -0.2485% | 1.853% | 0.824% | 4.933% | -2.895% | -4.204% | -21.94%
| Z-Score (Absolute) |  | 2.62 | 0.67 | 0.05 | 0.76 | 0.44 | 1.03 | NA | 3.57 | 0.34 | 0.84 | 0.09 | 0.59 | 0 | 0.27 | 0.55 | 0.15 | 1.75 | 1.3 | 1.81 | 8.73
| Analytical results for Dy in OREAS 554 (Certified Value 3.54 ppm) | [22 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
3
| 1 | 1 | 01: 3.54; 02: 4; 03: 3.5; 04: 3.5; 06: 3.8; 09: 3.64; 10: 3.76; 11: 3.46; 13: 2.97; 14: 2.5; 15: 3.91 | 1
| 1 | 2 | 01: 3.56; 02: 3.5; 03: 3.5; 04: 3.5; 06: 3.5; 09: 3.81; 10: 3.55; 11: 3.44; 13: 2.98; 14: 2.41; 15: 3.97 | 0
| 1 | 3 | 01: 3.51; 02: 4; 03: 3.5; 04: 3.5; 06: 3.2; 09: 3.64; 10: 3.27; 11: 3.58; 13: 2.89; 14: 2.35; 15: 3.94 | 16
| 1 | 4 | 01: 3.51; 02: 3.5; 03: 3.5; 04: 3.5; 06: 3.6; 09: 3.63; 10: 3.02; 11: 3.44; 13: 3.03; 14: 2.42; 15: 3.75 | 3.5352
| 1 | 5 | 01: 3.41; 02: 3.5; 03: 3.5; 04: 4; 06: 3.3; 09: 3.73; 10: 3.49; 11: 3.63; 13: 3; 14: 2.51; 15: 3.75 | 44
| 1 | 6 | 01: 3.47; 02: 3.5; 03: 3.5; 04: 3.5; 06: 3.6; 09: 3.82; 10: 3.48; 11: 3.3; 13: 3.01; 14: 2.42; 15: 3.88
| Mean |  | 01: 3.5; 02: 3.6667; 03: 3.5; 04: 3.5833; 06: 3.5; 09: 3.7117; 10: 3.4283; 11: 3.475; 13: 2.98; 14: 2.435; 15: 3.8667
| Median |  | 01: 3.51; 02: 3.5; 03: 3.5; 04: 3.5; 06: 3.55; 09: 3.685; 10: 3.485; 11: 3.45; 13: 2.99; 14: 2.42; 15: 3.895
| Std Dev. |  | 01: 0.053666; 02: 0.2582; 03: 0; 04: 0.20412; 06: 0.21909; 09: 0.087958; 10: 0.2542; 11: 0.11692; 13: 0.04899; 14: 0.060249; 15: 0.095219
| Rel.Std.Dev. |  | 01: 1.533%; 02: 7.042%; 03: 0%; 04: 5.696%; 06: 6.26%; 09: 2.37%; 10: 7.415%; 11: 3.365%; 13: 1.644%; 14: 2.474%; 15: 2.463%
| PDM3 |  | 01: -0.9959%; 02: 3.719%; 03: -0.9959%; 04: 1.361%; 06: -0.9959%; 09: 4.991%; 10: -3.023%; 11: -1.703%; 13: -15.71%; 14: -31.12%; 15: 9.376%
| Z-Score (Absolute) |  | 01: 0; 02: 1.35; 03: 0; 04: 0.67; 06: 0; 09: 1.71; 10: 0.58; 11: 0.2; 13: 4.21; 14: 8.62; 15: 2.97
| Analytical results for Er in OREAS 554 (Certified Value 1.99 ppm) | [22 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
3
| 1 | 1 | 01: 2.09; 02: 2; 03: 2; 04: 2; 06: 2.2; 09: 2.11; 10: 2.17; 11: 1.96; 13: 1.84; 14: 1; 15: 2.22 | 1
| 1 | 2 | 01: 1.97; 02: 2; 03: 2; 04: 2; 06: 2.6; 09: 2.14; 10: 2.11; 11: 2; 13: 1.81; 14: 1.12; 15: 1.87 | 0
| 1 | 3 | 01: 2.03; 02: 2; 03: 2; 04: 1.9; 06: 1.7; 09: 2.08; 10: 1.99; 11: 2.02; 13: 1.89; 14: 1.02; 15: 2.55 | 16
| 1 | 4 | 01: 1.99; 02: 2.5; 03: 2; 04: 1.9; 06: 1.8; 09: 2.06; 10: 2.04; 11: 1.99; 13: 1.84; 14: 1.08; 15: 2.48 | 1.9881
| 1 | 5 | 01: 2.01; 02: 2; 03: 2; 04: 2.4; 06: 1.8; 09: 2.13; 10: 2.03; 11: 2.11; 13: 1.8; 14: 1.12; 15: 2.38 | 45
| 1 | 6 | 01: 2.11; 02: 2; 03: 2; 04: 2.1; 06: 1.8; 09: 2.12; 10: 2.13; 11: 1.89; 13: 1.93; 14: 1.13; 15: 2.46
| Mean |  | 01: 2.0333; 02: 2.0833; 03: 2; 04: 2.05; 06: 1.9833; 09: 2.1067; 10: 2.0783; 11: 1.995; 13: 1.8517; 14: 1.0783; 15: 2.3267
| Median |  | 01: 2.02; 02: 2; 03: 2; 04: 2; 06: 1.8; 09: 2.115; 10: 2.075; 11: 1.995; 13: 1.84; 14: 1.1; 15: 2.42
| Std Dev. |  | 01: 0.055737; 02: 0.20412; 03: 0; 04: 0.18708; 06: 0.34881; 09: 0.030768; 10: 0.068823; 11: 0.072319; 13: 0.049565; 14: 0.056006; 15: 0.25057
| Rel.Std.Dev. |  | 01: 2.741%; 02: 9.798%; 03: 0%; 04: 9.126%; 06: 17.59%; 09: 1.461%; 10: 3.311%; 11: 3.625%; 13: 2.677%; 14: 5.194%; 15: 10.77%
| PDM3 |  | 01: 2.274%; 02: 4.789%; 03: 0.5973%; 04: 3.112%; 06: -0.241%; 09: 5.962%; 10: 4.537%; 11: 0.3458%; 13: -6.864%; 14: -45.76%; 15: 17.03%
| Z-Score (Absolute) |  | 01: 0.11; 02: 0.54; 03: NA; 04: 0.11; 06: 0.75; 09: 0.83; 10: 0.47; 11: 0.6; 13: 2.44; 14: 12.37; 15: 3.66
Results from laboratory 4 were removed due to their 1 ppm reading resolution.
| Analytical results for Eu in OREAS 554 (Certified Value 0.85 ppm) | [22 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 20; 10: 21; 11: 28; 13: 29; 14: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES; 13: PF*MS; 14: PF*OES | 2
2
| 1 | 1 | 01: 0.84; 02: 0.8; 03: 0.8; 04: 0.5; 06: 0.9; 09: 0.83; 10: 0.83; 11: 0.75; 13: 1.3; 14: 0.98 | 1
| 1 | 2 | 01: 0.77; 02: 0.8; 03: 0.8; 04: 0.5; 06: 1.2; 09: 0.78; 10: 0.83; 11: 0.75; 13: 1.2; 14: 1.19 | 0
| 1 | 3 | 01: 0.91; 02: 1; 03: 0.8; 04: 0.5; 06: 1.1; 09: 0.74; 10: 0.84; 11: 0.74; 13: 1.31; 14: 1.08 | 16
| 1 | 4 | 01: 0.77; 02: 0.8; 03: 0.8; 04: 0.5; 06: 1.2; 09: 0.66; 10: 0.82; 11: 0.76; 13: 1.21; 14: 1.07 | 0.84767
| 1 | 5 | 01: 0.84; 02: 1; 03: 1; 04: 0.5; 06: 1.2; 09: 0.73; 10: 0.81; 11: 0.72; 13: 1.24; 14: 1.1 | 46
| 1 | 6 | 01: 0.88; 02: 0.8; 03: 0.8; 04: 0.5; 06: 1.2; 09: 0.78; 10: 0.81; 11: 0.72; 13: 1.26; 14: 1.1
| Mean |  | 01: 0.835; 02: 0.86667; 03: 0.83333; 04: 0.5; 06: 1.1333; 09: 0.75333; 10: 0.82333; 11: 0.74; 13: 1.2533; 14: 1.0867
| Median |  | 01: 0.84; 02: 0.8; 03: 0.8; 04: 0.5; 06: 1.2; 09: 0.76; 10: 0.825; 11: 0.745; 13: 1.25; 14: 1.09
| Std Dev. |  | 01: 0.056833; 02: 0.10328; 03: 0.08165; 04: 0; 06: 0.12111; 09: 0.05785; 10: 0.012111; 11: 0.016733; 13: 0.045461; 14: 0.067429
| Rel.Std.Dev. |  | 01: 6.806%; 02: 11.92%; 03: 9.798%; 04: 0%; 06: 10.69%; 09: 7.679%; 10: 1.471%; 11: 2.261%; 13: 3.627%; 14: 6.205%
| PDM3 |  | 01: -1.494%; 02: 2.241%; 03: -1.691%; 04: -41.01%; 06: 33.7%; 09: -11.13%; 10: -2.871%; 11: -12.7%; 13: 47.86%; 14: 28.2%
| Z-Score (Absolute) |  | 01: 0.05; 02: NA; 03: NA; 04: NA; 06: NA; 09: 0.62; 10: 0.05; 11: 0.73; 13: 3.47; 14: 2.1
Results from laboratories 3, 4, 6 & 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 554 (Certified Value 1.82 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  | 2
3
| 1 | 1 | 1.9 | 1.74 | 1.78 | 1.87 | 1.8 | 1.89 | 1.83 | 1.89 | 1.777 | 1.77 | 1.763 | 1.868 | 1.88 | 1.786 | 1.79 | 1.7101 | 2.04 | 1.83 | 1.78 | 1.819 |  | 1
| 1 | 2 | 1.91 | 1.72 | 1.86 | 1.85 | 1.84 | 1.85 | 1.81 | 1.88 | 1.812 | 1.77 | 1.791 | 1.798 | 1.9234 | 1.786 | 1.82 | 1.8 | 2.08 | 1.83 | 1.74 | 1.84 |  | 0
| 1 | 3 | 1.93 | 1.76 | 1.83 | 1.88 | 1.81 | 1.84 | 1.82 | 1.8 | 1.791 | 1.8 | 1.777 | 1.819 | 1.8896 | 1.774 | 1.86 | 1.5653 | 2.04 | 1.89 | 1.75 | 1.833 |  | 16
| 1 | 4 | 1.85 | 1.71 | 1.86 | 1.84 | 1.84 | 1.87 | 1.82 | 1.89 | 1.791 | 1.8 | 1.77 | 1.819 | 1.7884 | 1.777 | 1.84 | 1.6505 | 2.04 | 1.82 | 1.77 | 1.791 |  | 1.8192
| 1 | 5 | 1.91 | 1.69 | 1.85 | 1.87 | 1.83 | 1.75 | 1.81 | 1.88 | 1.819 | 1.84 | 1.798 | 1.777 | 1.8621 | 1.758 | 1.84 | 1.6565 | 2.07 | 1.8 | 1.77 | 1.826 |  | 47
| 1 | 6 | 1.9 | 1.77 | 1.81 | 1.87 | 1.84 | 1.82 | 1.82 | 1.85 | 1.763 | 1.8 | 1.777 | 1.777 | 1.7841 | 1.778 | 1.88 | 1.5826 | 2.06 | 1.78 | 1.78 | 1.833
| Mean |  | 1.9 | 1.7317 | 1.8317 | 1.8633 | 1.8267 | 1.8367 | 1.8183 | 1.865 | 1.7922 | 1.7967 | 1.7793 | 1.8097 | 1.8546 | 1.7765 | 1.8383 | 1.6608 | 2.055 | 1.825 | 1.765 | 1.8237
| Median |  | 1.905 | 1.73 | 1.84 | 1.87 | 1.835 | 1.845 | 1.82 | 1.88 | 1.791 | 1.8 | 1.777 | 1.8085 | 1.8711 | 1.7775 | 1.84 | 1.6535 | 2.05 | 1.825 | 1.77 | 1.8295
| Std Dev. |  | 0.026833 | 0.030605 | 0.031885 | 0.015055 | 0.017512 | 0.048854 | 0.0075277 | 0.035071 | 0.020961 | 0.02582 | 0.013033 | 0.034197 | 0.056591 | 0.01031 | 0.031252 | 0.086217 | 0.017607 | 0.037283 | 0.016432 | 0.017523
| Rel.Std.Dev. |  | 1.412% | 1.767% | 1.741% | 0.808% | 0.9587% | 2.66% | 0.414% | 1.881% | 1.17% | 1.437% | 0.7325% | 1.89% | 3.051% | 0.5804% | 1.7% | 5.191% | 0.8568% | 2.043% | 0.931% | 0.9609%
| PDM3 |  | 4.442% | -4.812% | 0.6854% | 2.426% | 0.4105% | 0.9602% | -0.04757% | 2.518% | -1.486% | -1.239% | -2.191% | -0.524% | 1.946% | -2.347% | 1.052% | -8.705% | 12.96% | 0.3189% | -2.979% | 0.2456%
| Z-Score (Absolute) |  | 1.63 | 1.98 | 0.16 | 0.84 | 0.05 | 0.27 | 0.12 | 0.88 | 0.69 | 0.59 | 0.97 | 0.32 | 0.65 | 1.02 | 0.31 | 3.5 | 4.95 | 0.02 | 1.27 | 0.02
| Analytical results for Ga in OREAS 554 (Certified Value 13.5 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [7 empty] | 1
1
| 1 | 1 | 12 | 11 | 14 | < 100 | 12.1 | 12.4 | NA | 15.1441 | 11.9 | 11 | 15.5 | 14 | 12 | 15.2 | [7 empty] | 1
| 1 | 2 | 12 | 13 | 14 | < 100 | 15.3 | 12.4 | NA | 13.5444 | 12.2 | 11 | 16.1 | 13 | 13 | 14.5 | [7 empty] | 0
| 1 | 3 | 13 | 11 | 14 | < 100 | 15.2 | 12.3 | NA | 13.3605 | 11.2 | 11 | 16.2 | 13 | 12 | 14.4 | [7 empty] | 16
| 1 | 4 | 12 | 12 | 12 | < 100 | 14 | 12.5 | NA | 14.0172 | 11.9 | 11 | 16.7 | 14 | 13 | 15.9 | [7 empty] | 13.5184
| 1 | 5 | 12 | 12 | 12 | < 100 | 14.2 | 12.3 | NA | 13.9808 | 11.3 | 12 | 16 | 14 | 13 | 14.7 | [7 empty] | 48
| 1 | 6 | 13 | 14 | 12 | < 100 | 12.7 | 12.3 | NA | 14.0014 | 11.8 | 11 | 15.7 | 14 | 11 | 15.3
| Mean |  | 01: 12.3333; 02: 12.1667; 03: 13; 06: 13.9167; 09: 12.3667; 11: 14.0081; 13: 11.7167; 14: 11.1667; 15: 16.0333; 16: 13.6667; 18: 12.3333; 19: 15
| Median |  | 01: 12; 02: 12; 03: 13; 06: 14.1; 09: 12.35; 11: 13.9911; 13: 11.85; 14: 11; 15: 16.05; 16: 14; 18: 12.5; 19: 14.95
| Std Dev. |  | 01: 0.5164; 02: 1.169; 03: 1.0954; 06: 1.2983; 09: 0.08165; 11: 0.6206; 13: 0.38687; 14: 0.40825; 15: 0.41793; 16: 0.5164; 18: 0.8165; 19: 0.57271
| Rel.Std.Dev. |  | 01: 4.187%; 02: 9.609%; 03: 8.427%; 06: 9.329%; 09: 0.6602%; 11: 4.43%; 13: 3.302%; 14: 3.656%; 15: 2.607%; 16: 3.779%; 18: 6.62%; 19: 3.818%
| PDM3 |  | 01: -8.767%; 02: -9.999%; 03: -3.835%; 06: 2.946%; 09: -8.52%; 11: 3.622%; 13: -13.33%; 14: -17.4%; 15: 18.6%; 16: 1.096%; 18: -8.767%; 19: 10.96%
| Z-Score (Absolute) |  | NA | NA | 0.4 | 15.7 | 0 | 0.67 | NA | 0.04 | 0.96 | NA | 0.92 | NA | 0.69 | 0.47
Results from laboratories 1, 2, 21 & 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 554 (Certified Value 4.27 ppm) | [22 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
2
| 1 | 1 | 01: 4.1; 02: 4; 03: 4; 04: 5; 06: 4.9; 09: 4.96; 10: 4.85; 11: 3.92; 13: 4; 14: 3.88; 15: 4.3 | 1
| 1 | 2 | 01: 4.12; 02: 4; 03: 4; 04: 4; 06: 4.2; 09: 4.85; 10: 4.65; 11: 3.94; 13: 4; 14: 3.78; 15: 4.1 | 0
| 1 | 3 | 01: 4.1; 02: 4; 03: 4; 04: 5; 06: 4; 09: 4.99; 10: 4.38; 11: 3.91; 13: 4; 14: 3.56; 15: 4.88 | 16
| 1 | 4 | 01: 3.97; 02: 6; 03: 4; 04: 4; 06: 4.8; 09: 4.69; 10: 4.07; 11: 3.99; 13: 4; 14: 3.88; 15: 3.82 | 4.2692
| 1 | 5 | 01: 3.89; 02: 4; 03: 4; 04: 5; 06: 4.7; 09: 4.88; 10: 4.49; 11: 4.29; 13: 4; 14: 3.78; 15: 4.85 | 49
| 1 | 6 | 01: 3.79; 02: 4; 03: 4; 04: 4.5; 06: 4.2; 09: 4.6; 10: 4.6; 11: 3.78; 13: 4; 14: 3.65; 15: 4.6
| Mean |  | 01: 3.995; 02: 4.3333; 03: 4; 04: 4.5833; 06: 4.4667; 09: 4.8283; 10: 4.5067; 11: 3.9717; 13: 4; 14: 3.755; 15: 4.425
| Median |  | 01: 4.035; 02: 4; 03: 4; 04: 4.75; 06: 4.45; 09: 4.865; 10: 4.545; 11: 3.93; 13: 4; 14: 3.78; 15: 4.45
| Std Dev. |  | 01: 0.13517; 02: 0.8165; 03: 0; 04: 0.4916; 06: 0.37771; 09: 0.15355; 10: 0.26613; 11: 0.17081; 13: 0; 14: 0.12771; 15: 0.42557
| Rel.Std.Dev. |  | 01: 3.383%; 02: 18.84%; 03: 0%; 04: 10.73%; 06: 8.456%; 09: 3.18%; 10: 5.905%; 11: 4.301%; 13: 0%; 14: 3.401%; 15: 9.617%
| PDM3 |  | 01: -6.424%; 02: 1.501%; 03: -6.306%; 04: 7.357%; 06: 4.624%; 09: 13.1%; 10: 5.561%; 11: -6.97%; 13: -6.306%; 14: -12.05%; 15: 3.648%
| Z-Score (Absolute) |  | 01: 0.72; 02: NA; 03: NA; 04: NA; 06: 0.07; 09: 0.67; 10: 0.14; 11: 0.76; 13: NA; 14: 1.12; 15: 0
Results from laboratories 3, 4, 6 & 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 554 (Indicative Value 1.45 ppm) | [22 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 02; 03: 06; 04: 10; 06: 11; 09: 18; 10: 19; 11: 20; 13: 21; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 06: PF*MS; 09: PF*OES; 10: PF*MS; 11: PF*OES/MS; 13: PF*MS; 14: PF*MS; 15: PF*OES | 2
2
| 1 | 1 | 01: 1; 02: 2; 03: 3; 04: 2; 06: 1.5; 09: NA; 10: 1.5955; 11: < 5; 13: 1; 14: < 1; 15: 0.2 | 1
| 1 | 2 | 01: 1; 02: 2; 03: 3; 04: 2; 06: 1.3; 09: NA; 10: 1.3412; 11: < 5; 13: 1; 14: < 1; 15: 0.2 | 0
| 1 | 3 | 01: 2; 02: 1; 03: 3; 04: 1.7; 06: 1.4; 09: NA; 10: 1.4124; 11: < 5; 13: 1; 14: < 1; 15: 0.2 | 16
| 1 | 4 | 01: 2; 02: 1; 03: 2; 04: 1.9; 06: 1.5; 09: NA; 10: 1.4743; 11: < 5; 13: 1; 14: < 1; 15: 0.2 | 1.4542
| 1 | 5 | 01: 1; 02: 1; 03: 3; 04: 1.7; 06: 1.4; 09: NA; 10: 1.4853; 11: < 5; 13: 1; 14: < 1; 15: 0.2 | 21
| 1 | 6 | 01: 2; 02: 2; 03: 2; 04: 1.8; 06: 1.4; 09: NA; 10: 1.4433; 11: < 5; 13: 1; 14: < 1; 15: 0.2
| Mean |  | 01: 1.5; 02: 1.5; 03: 2.6667; 04: 1.85; 06: 1.4167; 10: 1.4587; 13: 1; 15: 0.2
| Median |  | 01: 1.5; 02: 1.5; 03: 3; 04: 1.85; 06: 1.4; 10: 1.4588; 13: 1; 15: 0.2
| Std Dev. |  | 01: 0.54772; 02: 0.54772; 03: 0.5164; 04: 0.13784; 06: 0.075277; 10: 0.084655; 13: 0; 15: 3.0405e-17
| Rel.Std.Dev. |  | 01: 36.51%; 02: 36.51%; 03: 19.36%; 04: 7.451%; 06: 5.314%; 10: 5.804%; 13: 0%; 15: 1.52e-14%
| PDM3 |  | 01: 3.148%; 02: 3.148%; 03: 83.37%; 04: 27.22%; 06: -2.583%; 10: 0.3063%; 13: -31.23%; 15: -86.25%
| Z-Score (Absolute) |  | 01: 0.03; 02: 0.03; 03: 1.88; 04: 0.59; 06: 0.1; 09: NA; 10: 0.03; 11: 1.62; 13: 0.76; 14: 1.55; 15: 2.03
| Analytical results for Hf in OREAS 554 (Certified Value 2.24 ppm) | [22 empty] | Certified
Hf | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 01; 02: 04; 03: 10; 04: 18; 06: 20; 09: 21; 10: 29; 11: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES; 06: PF*OES/MS; 09: PF*MS; 10: PF*MS; 11: PF*OES | 2
2
| 1 | 1 | 01: 2; 02: 2; 03: < 10; 04: NA; 06: < 5; 09: 2; 10: 2; 11: 3.3 | 1
| 1 | 2 | 01: 2; 02: 2; 03: < 10; 04: NA; 06: < 5; 09: 2; 10: 2; 11: 3 | 0
| 1 | 3 | 01: 2; 02: 2; 03: < 10; 04: NA; 06: < 5; 09: 2; 10: 2; 11: 2.9 | 16
| 1 | 4 | 01: 2; 02: 2; 03: < 10; 04: NA; 06: < 5; 09: 3; 10: 2; 11: 2.5 | 2.24
| 1 | 5 | 01: 2; 02: 2; 03: < 10; 04: NA; 06: < 5; 09: 3; 10: 2; 11: 2.8 | 50
| 1 | 6 | 01: 2; 02: 2; 03: < 10; 04: NA; 06: < 5; 09: 2; 10: 2; 11: 2.7
| Mean |  | 01: 2; 02: 2; 09: 2.3333; 10: 2; 11: 2.8667
| Median |  | 01: 2; 02: 2; 09: 2; 10: 2; 11: 2.85
| Std Dev. |  | 01: 0; 02: 0; 09: 0.5164; 10: 0; 11: 0.27325
| Rel.Std.Dev. |  | 01: 0%; 02: 0%; 09: 22.13%; 10: 0%; 11: 9.532%
| PDM3 |  | 01: -10.71%; 02: -10.71%; 09: 4.167%; 10: -10.71%; 11: 27.98%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67; 03: 5.39; 04: NA; 06: 0.34; 09: 0; 10: 0.67; 11: 1.08
| Analytical results for Hg in OREAS 554 (Indicative Value < 0.1 ppm) | [22 empty] | Indicative
Hg | Round | Replicate | 01: Lab; 02: Lab | 1
| No | No | 01: 18; 02: 30 | ppm
|  |  | 01: PF*OES; 02: PF*OES | 2
2
| 1 | 1 | 01: NA; 02: < 0.1 | 1
| 1 | 2 | 01: NA; 02: < 0.1 | 0
| 1 | 3 | 01: NA; 02: < 0.1 | 16
| 1 | 4 | 01: NA; 02: < 0.1 | < 0.1
| 1 | 5 | 01: NA; 02: < 0.1 | 22
| 1 | 6 | 01: NA; 02: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 02
| Analytical results for Ho in OREAS 554 (Certified Value 0.72 ppm) | [22 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
3
| 1 | 1 | 01: 0.74; 02: 0.6; 03: 0.6; 04: 0.6; 06: 0.5; 09: 0.76; 10: 0.77; 11: 0.68; 13: 0.6; 14: 0.5; 15: 0.79 | 1
| 1 | 2 | 01: 0.72; 02: 0.6; 03: 0.6; 04: 0.6; 06: 0.7; 09: 0.71; 10: 0.84; 11: 0.69; 13: 0.6; 14: 0.47; 15: 0.71 | 0
| 1 | 3 | 01: 0.71; 02: 0.6; 03: 0.8; 04: 0.7; 06: 0.7; 09: 0.73; 10: 0.67; 11: 0.72; 13: 0.59; 14: 0.51; 15: 0.85 | 16
| 1 | 4 | 01: 0.72; 02: 0.8; 03: 0.8; 04: 0.6; 06: 0.7; 09: 0.71; 10: 0.62; 11: 0.69; 13: 0.58; 14: 0.41; 15: 0.68 | 0.72467
| 1 | 5 | 01: 0.71; 02: 0.8; 03: 0.8; 04: 0.7; 06: 0.7; 09: 0.75; 10: 0.7; 11: 0.74; 13: 0.6; 14: 0.41; 15: 0.7 | 51
| 1 | 6 | 01: 0.71; 02: 0.6; 03: 0.6; 04: 0.7; 06: 0.7; 09: 0.73; 10: 0.71; 11: 0.67; 13: 0.62; 14: 0.4; 15: 0.81
| Mean |  | 01: 0.71833; 02: 0.66667; 03: 0.7; 04: 0.65; 06: 0.66667; 09: 0.73167; 10: 0.71833; 11: 0.69833; 13: 0.59833; 14: 0.45; 15: 0.75667
| Median |  | 01: 0.715; 02: 0.6; 03: 0.7; 04: 0.65; 06: 0.7; 09: 0.73; 10: 0.705; 11: 0.69; 13: 0.6; 14: 0.44; 15: 0.75
| Std Dev. |  | 01: 0.01169; 02: 0.10328; 03: 0.10954; 04: 0.054772; 06: 0.08165; 09: 0.020412; 10: 0.077309; 11: 0.026394; 13: 0.013292; 14: 0.049396; 15: 0.069186
| Rel.Std.Dev. |  | 01: 1.627%; 02: 15.49%; 03: 15.65%; 04: 8.427%; 06: 12.25%; 09: 2.79%; 10: 10.76%; 11: 3.78%; 13: 2.221%; 14: 10.98%; 15: 9.143%
| PDM3 |  | 01: -0.874%; 02: -8.004%; 03: -3.404%; 04: -10.3%; 06: -8.004%; 09: 0.966%; 10: -0.874%; 11: -3.634%; 13: -17.43%; 14: -37.9%; 15: 4.416%
| Z-Score (Absolute) |  | 01: 0; 02: NA; 03: NA; 04: NA; 06: NA; 09: 0.45; 10: 0; 11: 0.67; 13: 4.05; 14: 9.05; 15: 1.29
Results from laboratories 3, 4, 6 & 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 554 (Indicative Value < 0.2 ppm) | [22 empty] | Indicative
In | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 06: 06; 09: 10; 10: 11; 11: 18; 13: 20; 14: 21; 15: 23; 16: 29 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*MS; 06: PF*OES/MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES; 13: PF*OES/MS; 14: PF*MS; 15: PF*MS; 16: PF*MS | 2
2
| 1 | 1 | 01: < 0.2; 02: < 0.1; 03: 0.2; 04: 0.2; 06: < 0.5; 09: < 0.2; 10: < 0.2; 11: NA; 13: < 0.5; 14: < 0.2; 15: 0.25; 16: < 0.2 | 1
| 1 | 2 | 01: < 0.2; 02: 0.1; 03: 0.2; 04: 0.2; 06: < 0.5; 09: < 0.2; 10: < 0.2; 11: NA; 13: < 0.5; 14: < 0.2; 15: 0.23; 16: < 0.2 | 0
| 1 | 3 | 01: < 0.2; 02: < 0.1; 03: < 0.2; 04: 0.2; 06: < 0.5; 09: < 0.2; 10: < 0.2; 11: NA; 13: < 0.5; 14: < 0.2; 15: 0.22; 16: < 0.2 | 16
| 1 | 4 | 01: < 0.2; 02: < 0.1; 03: 0.2; 04: 0.2; 06: < 0.5; 09: < 0.2; 10: < 0.2; 11: NA; 13: < 0.5; 14: < 0.2; 15: 0.28; 16: < 0.2 | < 0.2
| 1 | 5 | 01: < 0.2; 02: < 0.1; 03: 0.2; 04: 0.2; 06: < 0.5; 09: < 0.2; 10: < 0.2; 11: NA; 13: < 0.5; 14: < 0.2; 15: 0.26; 16: < 0.2 | 23
| 1 | 6 | 01: < 0.2; 02: 0.2; 03: 0.2; 04: 0.2; 06: < 0.5; 09: < 0.2; 10: < 0.2; 11: NA; 13: < 0.5; 14: < 0.2; 15: 0.27; 16: < 0.2
| Mean |  | 02: 0.15; 03: 0.2; 04: 0.2; 15: 0.25167
| Median |  | 02: 0.15; 03: 0.2; 04: 0.2; 15: 0.255
| Std Dev. |  | 02: 0.070711; 03: 0; 04: 3.0405e-17; 15: 0.023166
| Rel.Std.Dev. |  | 02: 47.14%; 03: 0%; 04: 1.52e-14%; 15: 9.205%
PDM3
| Z-Score (Absolute) |  | 01: 0; 02: 0.67; 03: 3.37; 04: 4.05; 06: 6.07; 09: 0; 10: 0; 11: NA; 13: 6.07; 14: 0; 15: 6.14; 16: 0
| Analytical results for K in OREAS 554 (Certified Value 1.69 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  |  | 2
3
| 1 | 1 | 1.6 | 1.71 | 1.7 | 1.7 | 1.7 | 1.86 | 1.7 | 1.6 | 1.6 | 1.66 | 1.727 | NA | 1.8977 | 1.67 | 2.0414 | 1.99 | 0.9 | 1.68 | 1.71 |  |  | 1
| 1 | 2 | 1.7 | 1.67 | 1.75 | 1.7 | 1.7 | 1.73 | 1.7 | 1.6 | 1.6 | 1.677 | 1.694 | NA | 1.8562 | 1.74 | 2.1326 | 2.01 | 1.7 | 1.64 | 1.735 |  |  | 0
| 1 | 3 | 1.6 | 1.74 | 1.7 | 1.7 | 1.7 | 1.82 | 1.7 | 1.7 | 1.6 | 1.702 | 1.702 | NA | 1.818 | 1.79 | 2.0793 | 1.95 | 1.7 | 1.63 | 1.71 |  |  | 16
| 1 | 4 | 1.6 | 1.67 | 1.65 | 1.7 | 1.7 | 1.71 | 1.7 | 1.7 | 1.5 | 1.71 | 1.677 | NA | 1.8214 | 1.76 | 2.0133 | 1.95 | 1.6 | 1.64 | 1.677 |  |  | 1.6921
| 1 | 5 | 1.6 | 1.69 | 1.7 | 1.7 | 1.7 | 1.68 | 1.7 | 1.6 | 1.6 | 1.694 | 1.677 | NA | 1.7848 | 1.76 | 1.9179 | 1.96 | 1.6 | 1.65 | 1.71 |  |  | 52
| 1 | 6 | 1.6 | 1.73 | 1.75 | 1.7 | 1.7 | 1.68 | 1.7 | 1.6 | 1.6 | 1.702 | 1.702 | NA | 1.7898 | 1.81 | 1.9704 | 2 | 1.7 | 1.66 | 1.76
| Mean |  | 1.6167 | 1.7017 | 1.7083 | 1.7 | 1.7 | 1.7467 | 1.7 | 1.6333 | 1.5833 | 1.6908 | 1.6965 |  | 1.828 | 1.755 | 2.0258 | 1.9767 | 1.5333 | 1.65 | 1.717
| Median |  | 1.6 | 1.7 | 1.7 | 1.7 | 1.7 | 1.72 | 1.7 | 1.6 | 1.6 | 1.698 | 1.698 |  | 1.8197 | 1.76 | 2.0274 | 1.975 | 1.65 | 1.645 | 1.71
| Std Dev. |  | 0.040825 | 0.029944 | 0.037639 | 2.4324e-16 | 2.4324e-16 | 0.075807 | 2.4324e-16 | 0.05164 | 0.040825 | 0.018809 | 0.018748 |  | 0.042714 | 0.048477 | 0.076655 | 0.026583 | 0.31411 | 0.017889 | 0.028
| Rel.Std.Dev. |  | 2.525% | 1.76% | 2.203% | 1.431e-14% | 1.431e-14% | 4.34% | 1.431e-14% | 3.162% | 2.578% | 1.112% | 1.105% |  | 2.337% | 2.762% | 3.784% | 1.345% | 20.49% | 1.084% | 1.631%
| PDM3 |  | -4.456% | 0.5673% | 0.9613% | 0.4688% | 0.4688% | 3.227% | 0.4688% | -3.471% | -6.426% | -0.07298% | 0.2619% |  | 8.032% | 3.719% | 19.72% | 16.82% | -9.381% | -2.486% | 1.473%
| Z-Score (Absolute) |  | 1.16 | 0.02 | 0.12 | 0 | 0 | 0.65 | 0 | 0.93 | 1.63 | 0.13 | 0.05 | NA | 1.79 | 0.77 | 4.55 | 3.86 | 2.33 | 0.7 | 0.24
| Analytical results for La in OREAS 554 (Certified Value 35.2 ppm) | [22 empty] | Certified
La | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 18; 11: 20; 13: 21; 14: 28; 15: 29; 16: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: PF*OES/MS; 13: PF*MS; 14: PF*OES; 15: PF*MS; 16: PF*OES | 1
2
| 1 | 1 | 01: 36.7; 02: 36.5; 03: 33; 04: 36; 06: 36.1; 09: 36.7; 10: NA; 11: 37.4; 13: 30.1; 14: 31.99; 15: 26.3; 16: 35 | 1
| 1 | 2 | 01: 36.5; 02: 38; 03: 32; 04: 35; 06: 35.6; 09: 36.1; 10: NA; 11: 35.2; 13: 29.7; 14: 32.47; 15: 25.9; 16: 34 | 0
| 1 | 3 | 01: 36; 02: 37; 03: 32.5; 04: 37; 06: 37; 09: 36; 10: NA; 11: 33.3; 13: 29.7; 14: 33.17; 15: 26.5; 16: 38 | 16
| 1 | 4 | 01: 35.4; 02: 39; 03: 32.5; 04: 36; 06: 35; 09: 35.9; 10: NA; 11: 30.8; 13: 29.8; 14: 32.47; 15: 26.9; 16: 38 | 35.1863
| 1 | 5 | 01: 36; 02: 39; 03: 32.5; 04: 35; 06: 34.7; 09: 35.9; 10: NA; 11: 33.7; 13: 31.7; 14: 32.55; 15: 25.1; 16: 35 | 53
| 1 | 6 | 01: 36.1; 02: 37; 03: 32.5; 04: 35; 06: 36.3; 09: 35.9; 10: NA; 11: 34.8; 13: 29.2; 14: 31.81; 15: 25.6; 16: 37
| Mean |  | 01: 36.1167; 02: 37.75; 03: 32.5; 04: 35.6667; 06: 35.7833; 09: 36.0833; 11: 34.2; 13: 30.0333; 14: 32.41; 15: 26.05; 16: 36.1667
| Median |  | 01: 36.05; 02: 37.5; 03: 32.5; 04: 35.5; 06: 35.85; 09: 35.95; 11: 34.25; 13: 29.75; 14: 32.47; 15: 26.1; 16: 36
| Std Dev. |  | 01: 0.4535; 02: 1.084; 03: 0.31623; 04: 0.8165; 06: 0.85654; 09: 0.31252; 11: 2.2009; 13: 0.86641; 14: 0.47766; 15: 0.65038; 16: 1.7224
| Rel.Std.Dev. |  | 01: 1.256%; 02: 2.871%; 03: 0.973%; 04: 2.289%; 06: 2.394%; 09: 0.8661%; 11: 6.435%; 13: 2.885%; 14: 1.474%; 15: 2.497%; 16: 4.762%
| PDM3 |  | 01: 2.644%; 02: 7.286%; 03: -7.634%; 04: 1.365%; 06: 1.697%; 09: 2.549%; 11: -2.803%; 13: -14.64%; 14: -7.89%; 15: -25.97%; 16: 2.786%
| Z-Score (Absolute) |  | 01: 0.21; 02: 0.96; 03: 1.46; 04: 0; 06: 0.05; 09: 0.19; 10: NA; 11: 0.67; 13: 2.59; 14: 1.5; 15: 4.42; 16: 0.23
| Analytical results for Li in OREAS 554 (Certified Value 34.2 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 28 | 29 | [6 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | [6 empty] | 1
1
| 1 | 1 | 30 | 42 | 35 | 40 | 50 | 50 | 38 | 40 | NA | 21.8749 | 33 | 31 | 41 | 32 | 30 | [6 empty] | 1
| 1 | 2 | 28 | 43 | 35 | 40 | 50 | 50 | 38 | 41 | NA | 19.9333 | 34 | 31 | 41 | 34 | 31 | [6 empty] | 0
| 1 | 3 | 32 | 50 | 35 | 40 | 50 | 50 | 37 | 42 | NA | 20.2281 | 35 | 32 | 42 | 30 | 33 | [6 empty] | 16
| 1 | 4 | 32 | 39 | 36 | 40 | 60 | 40 | 38 | 42 | NA | 20.6545 | 35 | 31 | 41 | 29 | 32 | [6 empty] | 34.2376
| 1 | 5 | 29 | 47 | 36 | 40 | 50 | 30 | 50 | 41 | NA | 20.8753 | 34 | 32 | 42 | 31 | 29 | [6 empty] | 54
| 1 | 6 | 30 | 37 | 35 | 40 | 50 | 40 | 38 | 41 | NA | 20.3145 | 34 | 31 | 41 | 31 | 28
| Mean |  | 30.1667 | 43 | 35.3333 | 40 | 51.6667 | 43.3333 | 39.8333 | 41.1667 |  | 20.6468 | 34.1667 | 31.3333 | 41.3333 | 31.1667 | 30.5
| Median |  | 30 | 42.5 | 35 | 40 | 50 | 45 | 38 | 41 |  | 20.4845 | 34 | 31 | 41 | 31 | 30.5
| Std Dev. |  | 1.6021 | 4.858 | 0.5164 | 0 | 4.0825 | 8.165 | 4.9967 | 0.75277 |  | 0.68634 | 0.75277 | 0.5164 | 0.5164 | 1.7224 | 1.8708
| Rel.Std.Dev. |  | 5.311% | 11.3% | 1.462% | 0% | 7.902% | 18.84% | 12.54% | 1.829% |  | 3.324% | 2.203% | 1.648% | 1.249% | 5.526% | 6.134%
| PDM3 |  | -11.89% | 25.59% | 3.2% | 16.83% | 50.91% | 26.57% | 16.34% | 20.24% |  | -39.7% | -0.2071% | -8.483% | 20.73% | -8.969% | -10.92%
| Z-Score (Absolute) |  | 0.67 | 1.49 | 0.2 | NA | NA | NA | 0.96 | 1.18 | NA | 2.28 | 0 | 0.48 | 1.21 | 0.51 | 0.62
Results from laboratories 4, 6 & 9 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 554 (Certified Value 0.29 ppm) | [22 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 11; 09: 20; 10: 21; 11: 28; 13: 29; 14: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES; 13: PF*MS; 14: PF*OES | 2
2
| 1 | 1 | 01: 0.31; 02: 0.4; 03: 0.2; 04: 0.6; 06: 0.31; 09: < 0.5; 10: 0.28; 11: 0.29; 13: 0.3; 14: 0.36 | 1
| 1 | 2 | 01: 0.27; 02: 0.4; 03: 0.2; 04: 0.4; 06: 0.32; 09: < 0.5; 10: 0.28; 11: 0.29; 13: 0.25; 14: 0.35 | 0
| 1 | 3 | 01: 0.29; 02: 0.4; 03: 0.4; 04: 0.4; 06: 0.32; 09: < 0.5; 10: 0.27; 11: 0.29; 13: 0.26; 14: 0.34 | 16
| 1 | 4 | 01: 0.26; 02: 0.4; 03: 0.2; 04: 0.4; 06: 0.33; 09: < 0.5; 10: 0.28; 11: 0.29; 13: 0.24; 14: 0.37 | 0.2945
| 1 | 5 | 01: 0.26; 02: 0.2; 03: 0.2; 04: 0.6; 06: 0.32; 09: < 0.5; 10: 0.29; 11: 0.28; 13: 0.26; 14: 0.37 | 55
| 1 | 6 | 01: 0.29; 02: 0.2; 03: 0.2; 04: 0.2; 06: 0.31; 09: < 0.5; 10: 0.27; 11: 0.29; 13: 0.25; 14: 0.31
| Mean |  | 01: 0.28; 02: 0.33333; 03: 0.23333; 04: 0.43333; 06: 0.31833; 10: 0.27833; 11: 0.28833; 13: 0.26; 14: 0.35
| Median |  | 01: 0.28; 02: 0.4; 03: 0.2; 04: 0.4; 06: 0.32; 10: 0.28; 11: 0.29; 13: 0.255; 14: 0.355
| Std Dev. |  | 01: 0.02; 02: 0.10328; 03: 0.08165; 04: 0.15055; 06: 0.0075277; 10: 0.0075277; 11: 0.0040825; 13: 0.020976; 14: 0.022804
| Rel.Std.Dev. |  | 01: 7.143%; 02: 30.98%; 03: 34.99%; 04: 34.74%; 06: 2.365%; 10: 2.705%; 11: 1.416%; 13: 8.068%; 14: 6.515%
| PDM3 |  | 01: -4.924%; 02: 13.19%; 03: -20.77%; 04: 47.14%; 06: 8.093%; 10: -5.49%; 11: -2.094%; 13: -11.71%; 14: 18.85%
| Z-Score (Absolute) |  | 01: 0; 02: NA; 03: NA; 04: NA; 06: 1.29; 09: 1.01; 10: 0.06; 11: 0.28; 13: 0.67; 14: 2.36
Results from laboratories 3, 4 & 6 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 554 (Certified Value 3.14 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  | 2
3
| 1 | 1 | 3.17 | 3.15 | 3.09 | 3.05 | 3.24 | 3.26 | 3.13 | 3.29 | 3.172 | 3.07 | 3.142 | 3.232 | 3.1568 | 3.236 | 2.94 | 3.1214 | 3.47 | 2.04 | 2.85 | 3.106 |  | 1
| 1 | 2 | 3.15 | 3.14 | 3.07 | 3.08 | 3.23 | 3.11 | 3.12 | 3.29 | 3.142 | 3.142 | 3.142 | 3.136 | 3.1128 | 3.2077 | 3.15 | 3.0915 | 3.45 | 3.18 | 2.95 | 3.124 |  | 0
| 1 | 3 | 3.19 | 3.2 | 3.1 | 3.1 | 3.195 | 3.14 | 3.13 | 3.25 | 3.172 | 3.021 | 3.142 | 3.148 | 3.1403 | 3.1974 | 3.21 | 3.0207 | 3.47 | 3.22 | 2.98 | 3.082 |  | 16
| 1 | 4 | 3.1 | 3.08 | 3.1 | 3.09 | 3.25 | 3.22 | 3.14 | 3.26 | 3.148 | 3.021 | 3.142 | 3.172 | 3.1224 | 3.2131 | 3.12 | 3.0665 | 3.44 | 3.01 | 3.03 | 3.058 |  | 3.1429
| 1 | 5 | 3.15 | 3.12 | 3.15 | 3.08 | 3.2 | 3.01 | 3.09 | 3.22 | 3.16 | 3.148 | 3.148 | 3.178 | 3.1106 | 3.1987 | 3.19 | 2.7845 | 3.5 | 3.07 | 3.11 | 3.076 |  | 56
| 1 | 6 | 3.12 | 3.22 | 3.12 | 3.12 | 3.22 | 3.03 | 3.12 | 3.24 | 3.142 | 3.1 | 3.148 | 3.172 | 3.0619 | 3.1908 | 3.18 | 3.0742 | 3.49 | 3.16 | 3.06 | 3.142
| Mean |  | 3.1467 | 3.1517 | 3.105 | 3.0867 | 3.2225 | 3.1283 | 3.1217 | 3.2583 | 3.156 | 3.0837 | 3.144 | 3.173 | 3.1175 | 3.2073 | 3.1317 | 3.0265 | 3.47 | 2.9467 | 2.9967 | 3.098
| Median |  | 3.15 | 3.145 | 3.1 | 3.085 | 3.225 | 3.125 | 3.125 | 3.255 | 3.154 | 3.085 | 3.142 | 3.172 | 3.1176 | 3.2032 | 3.165 | 3.0703 | 3.47 | 3.115 | 3.005 | 3.094
| Std Dev. |  | 0.03266 | 0.051543 | 0.027386 | 0.023381 | 0.021852 | 0.099883 | 0.017224 | 0.027869 | 0.014029 | 0.056302 | 0.0030984 | 0.033172 | 0.032402 | 0.016149 | 0.099079 | 0.12305 | 0.022804 | 0.45076 | 0.091579 | 0.031673
| Rel.Std.Dev. |  | 1.038% | 1.635% | 0.882% | 0.7575% | 0.6781% | 3.193% | 0.5518% | 0.8553% | 0.4445% | 1.826% | 0.09855% | 1.045% | 1.039% | 0.5035% | 3.164% | 4.066% | 0.6572% | 15.3% | 3.056% | 1.022%
| PDM3 |  | 0.1185% | 0.2776% | -1.207% | -1.791% | 2.531% | -0.4648% | -0.677% | 3.671% | 0.4154% | -1.886% | 0.03362% | 0.9563% | -0.8106% | 2.047% | -0.3588% | -3.706% | 10.41% | -6.245% | -4.654% | -1.43%
| Z-Score (Absolute) |  | 0.3 | 0.39 | 0.45 | 0.78 | 1.65 | 0.03 | 0.15 | 2.3 | 0.47 | 0.83 | 0.25 | 0.77 | 0.22 | 1.38 | 0.03 | 1.85 | 6.08 | 3.28 | 2.39 | 0.58
| Analytical results for Mn in OREAS 554 (Certified Value 0.11 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  | 3
3
| 1 | 1 | 0.1057 | 0.1121 | 0.112 | 0.108 | 0.115 | 0.12 | 0.112 | 0.13 | 0.108 | 0.101 | 0.116 | 0.116 | NA | 0.1061 | 0.108 | 0.109 | 0.117 | 0.1035 | 0.1038 | 0.108 |  | 1
| 1 | 2 | 0.1031 | 0.1109 | 0.112 | 0.11 | 0.115 | 0.11 | 0.112 | 0.12 | 0.108 | 0.108 | 0.116 | 0.108 | NA | 0.1053 | 0.113 | 0.1128 | 0.116 | 0.1076 | 0.1027 | 0.108 |  | 0
| 1 | 3 | 0.1095 | 0.1129 | 0.11 | 0.112 | 0.115 | 0.11 | 0.11 | 0.12 | 0.108 | 0.101 | 0.116 | 0.108 | NA | 0.1046 | 0.118 | 0.106 | 0.114 | 0.1078 | 0.1052 | 0.108 |  | 16
| 1 | 4 | 0.0995 | 0.1096 | 0.114 | 0.11 | 0.115 | 0.11 | 0.114 | 0.12 | 0.108 | 0.101 | 0.108 | 0.108 | NA | 0.103 | 0.115 | 0.1278 | 0.113 | 0.1012 | 0.1064 | 0.101 |  | 0.1103
| 1 | 5 | 0.1046 | 0.1137 | 0.114 | 0.109 | 0.115 | 0.11 | 0.111 | 0.12 | 0.108 | 0.108 | 0.116 | 0.116 | NA | 0.1053 | 0.113 | 0.111 | 0.116 | 0.1036 | 0.1038 | 0.108 |  | 57
| 1 | 6 | 0.1093 | 0.1147 | 0.11 | 0.117 | 0.115 | 0.1 | 0.112 | 0.12 | 0.108 | 0.108 | 0.116 | 0.108 | NA | 0.103 | 0.116 | 0.1134 | 0.116 | 0.1072 | 0.1028 | 0.108
| Mean |  | 0.10528 | 0.11232 | 0.112 | 0.111 | 0.115 | 0.11 | 0.11183 | 0.12167 | 0.108 | 0.1045 | 0.11467 | 0.11067 |  | 0.10455 | 0.11383 | 0.11333 | 0.11533 | 0.10515 | 0.10412 | 0.10683
| Median |  | 0.10515 | 0.1125 | 0.112 | 0.11 | 0.115 | 0.11 | 0.112 | 0.12 | 0.108 | 0.1045 | 0.116 | 0.108 |  | 0.10495 | 0.114 | 0.1119 | 0.116 | 0.1054 | 0.1038 | 0.108
| Std Dev. |  | 0.0038149 | 0.0018638 | 0.0017889 | 0.0032249 | 0 | 0.0063246 | 0.0013292 | 0.0040825 | 0 | 0.0038341 | 0.003266 | 0.0041312 |  | 0.0012911 | 0.0034303 | 0.0075838 | 0.0015055 | 0.0027552 | 0.0014372 | 0.0028577
| Rel.Std.Dev. |  | 3.623% | 1.659% | 1.597% | 2.905% | 0% | 5.75% | 1.189% | 3.355% | 0% | 3.669% | 2.848% | 3.733% |  | 1.235% | 3.013% | 6.692% | 1.305% | 2.62% | 1.38% | 2.675%
| PDM3 |  | -4.546% | 1.831% | 1.544% | 0.637% | 4.264% | -0.2696% | 1.393% | 10.31% | -2.083% | -5.256% | 3.961% | 0.3348% |  | -5.211% | 3.206% | 2.753% | 4.566% | -4.667% | -5.604% | -3.141%
| Z-Score (Absolute) |  | 1 | 0.23 | 0.17 | 0 | 0.7 | 0.18 | 0.14 | 1.86 | 0.45 | 1.12 | 0.67 | 0 | NA | 1.12 | 0.49 | 0.41 | 0.75 | 1.02 | 1.2 | 0.67
| Analytical results for Mo in OREAS 554 (Indicative Value 4.18 ppm) | [22 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 29 | 30 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | PF*OES | [7 empty] | 2
2
| 1 | 1 | 3 | 5 | < 5 | < 5 | 5 | 5 | 3.6 | NA | 2.427 | < 5 | 3 | 5 | < 2 | 5.7 | [7 empty] | 1
| 1 | 2 | 3 | 8 | < 5 | < 5 | < 5 | 7 | 3.8 | NA | 2.2152 | < 5 | 3 | 5 | 2 | 5.7 | [7 empty] | 0
| 1 | 3 | 4 | 5 | < 5 | < 5 | < 5 | 3 | 3.8 | NA | 2.4865 | < 5 | 3 | 5 | < 2 | 6.9 | [7 empty] | 16
| 1 | 4 | 3 | 10 | < 5 | < 5 | < 5 | 4 | 3.8 | NA | 2.4621 | < 5 | 3 | 5 | 2 | 6.9 | [7 empty] | 4.1837
| 1 | 5 | 3 | 6 | < 5 | < 5 | < 5 | 5 | 3.7 | NA | 2.4868 | < 5 | 4 | 5 | 2 | 6.6 | [7 empty] | 19
| 1 | 6 | 4 | 4 | 5 | < 5 | < 5 | 4 | 3.5 | NA | 2.5729 | < 5 | 3 | 5 | 2 | 4.6
| Mean |  | 01: 3.3333; 02: 6.3333; 03: 5; 06: 5; 09: 4.6667; 10: 3.7; 13: 2.4417; 15: 3.1667; 16: 5; 18: 2; 19: 6.0667
| Median |  | 01: 3; 02: 5.5; 03: 5; 06: 5; 09: 4.5; 10: 3.75; 13: 2.4743; 15: 3; 16: 5; 18: 2; 19: 6.15
| Std Dev. |  | 01: 0.5164; 02: 2.2509; 09: 1.3663; 10: 0.12649; 13: 0.12098; 15: 0.40825; 16: 0; 18: 0; 19: 0.9048
| Rel.Std.Dev. |  | 01: 15.49%; 02: 35.54%; 09: 29.28%; 10: 3.419%; 13: 4.955%; 15: 12.89%; 16: 0%; 18: 0%; 19: 14.91%
| PDM3 |  | 01: -20.33%; 02: 51.38%; 03: 19.51%; 06: 19.51%; 09: 11.54%; 10: -11.56%; 13: -41.64%; 15: -24.31%; 16: 19.51%; 18: -52.2%; 19: 45.01%
| Z-Score (Absolute) |  | 0.17 | 3.2 | 0.25 | 0.67 | 0.25 | 1.52 | 0.54 | NA | 0.73 | 0.67 | 0 | 1.85 | 1.52 | 2.93
| Analytical results for Nb in OREAS 554 (Indicative Value 7.81 ppm) | [22 empty] | Indicative
Nb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 18; 10: 19; 11: 20; 13: 21; 14: 28; 15: 29; 16: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*OES; 10: PF*MS; 11: PF*OES/MS; 13: PF*MS; 14: PF*OES; 15: PF*MS; 16: PF*OES | 2
2
| 1 | 1 | 01: 7; 02: 5; 03: 10; 04: 30; 06: 9; 09: NA; 10: 9.9345; 11: 9; 13: < 10; 14: 13; 15: 4; 16: 7 | 1
| 1 | 2 | 01: 7; 02: 5; 03: 10; 04: 30; 06: 8.9; 09: NA; 10: 9.0421; 11: 10; 13: < 10; 14: 16; 15: 4; 16: 7 | 0
| 1 | 3 | 01: 7; 02: 10; 03: 10; 04: 20; 06: 9.2; 09: NA; 10: 9.8353; 11: 8; 13: < 10; 14: 15; 15: 3; 16: 7 | 16
| 1 | 4 | 01: 7; 02: 10; 03: 10; 04: 25; 06: 8.5; 09: NA; 10: 10.4172; 11: 8; 13: < 10; 14: 11; 15: 4; 16: 7 | 7.8081
| 1 | 5 | 01: 6; 02: 5; 03: 10; 04: 20; 06: 9.5; 09: NA; 10: 10.6093; 11: 8; 13: < 10; 14: 22; 15: 3; 16: 7 | 20
| 1 | 6 | 01: 7; 02: 10; 03: 10; 04: 20; 06: 9.1; 09: NA; 10: 9.7499; 11: 9; 13: < 10; 14: 15; 15: 3; 16: 7
| Mean |  | 01: 6.8333; 02: 7.5; 03: 10; 04: 24.1667; 06: 9.0333; 10: 9.9314; 11: 8.6667; 14: 15.3333; 15: 3.5; 16: 7
| Median |  | 01: 7; 02: 7.5; 03: 10; 04: 22.5; 06: 9.05; 10: 9.8849; 11: 8.5; 14: 15; 15: 3.5; 16: 7
| Std Dev. |  | 01: 0.40825; 02: 2.7386; 03: 0; 04: 4.916; 06: 0.33267; 10: 0.55291; 11: 0.8165; 14: 3.7238; 15: 0.54772; 16: 0
| Rel.Std.Dev. |  | 01: 5.974%; 02: 36.51%; 03: 0%; 04: 20.34%; 06: 3.683%; 10: 5.567%; 11: 9.421%; 14: 24.29%; 15: 15.65%; 16: 0%
| PDM3 |  | 01: -12.48%; 02: -3.946%; 03: 28.07%; 04: 209.5%; 06: 15.69%; 10: 27.19%; 11: 11%; 14: 96.38%; 15: -55.17%; 16: -10.35%
| Z-Score (Absolute) |  | 01: 0.74; 02: 0.47; 03: 0.54; 04: 6.27; 06: 0.15; 09: NA; 10: 0.51; 11: 0; 13: 1.48; 14: 2.7; 15: 2.09; 16: 0.67
| Analytical results for Nd in OREAS 554 (Certified Value 28.2 ppm) | [22 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 1
2
| 1 | 1 | 01: 28.8; 02: 28.5; 03: 28; 04: 28; 06: 28.6; 09: 29.1; 10: 30.8; 11: 25; 13: 25; 14: 21.3; 15: 29.1 | 1
| 1 | 2 | 01: 28.3; 02: 28.5; 03: 28.5; 04: 28; 06: 28.8; 09: 28.8; 10: 28.7; 11: 24.7; 13: 25.4; 14: 21.8; 15: 28.7 | 0
| 1 | 3 | 01: 27.4; 02: 27.5; 03: 28; 04: 28; 06: 28.9; 09: 29; 10: 26.9; 11: 24.5; 13: 26.1; 14: 21.1; 15: 28.9 | 16
| 1 | 4 | 01: 27.9; 02: 29; 03: 28.5; 04: 28; 06: 32; 09: 29; 10: 25.2; 11: 25; 13: 25.4; 14: 22; 15: 32.5 | 28.1822
| 1 | 5 | 01: 27.8; 02: 27.5; 03: 29; 04: 28; 06: 29.2; 09: 28.8; 10: 27.1; 11: 27.2; 13: 26.4; 14: 22; 15: 31.3 | 58
| 1 | 6 | 01: 28.3; 02: 29; 03: 29; 04: 27; 06: 29.9; 09: 28.4; 10: 28.5; 11: 24.4; 13: 25.5; 14: 21; 15: 29.3
| Mean |  | 01: 28.0833; 02: 28.3333; 03: 28.5; 04: 27.8333; 06: 29.5667; 09: 28.85; 10: 27.8667; 11: 25.1333; 13: 25.6333; 14: 21.5333; 15: 29.9667
| Median |  | 01: 28.1; 02: 28.5; 03: 28.5; 04: 28; 06: 29.05; 09: 28.9; 10: 27.8; 11: 24.85; 13: 25.45; 14: 21.55; 15: 29.2
| Std Dev. |  | 01: 0.48751; 02: 0.68313; 03: 0.44721; 04: 0.40825; 06: 1.2754; 09: 0.251; 10: 1.9149; 11: 1.0424; 13: 0.5164; 14: 0.45461; 15: 1.5578
| Rel.Std.Dev. |  | 01: 1.736%; 02: 2.411%; 03: 1.569%; 04: 1.467%; 06: 4.314%; 09: 0.87%; 10: 6.871%; 11: 4.148%; 13: 2.015%; 14: 2.111%; 15: 5.198%
| PDM3 |  | 01: -0.3509%; 02: 0.5362%; 03: 1.128%; 04: -1.238%; 06: 4.912%; 09: 2.37%; 10: -1.12%; 11: -10.82%; 13: -9.044%; 14: -23.59%; 15: 6.332%
| Z-Score (Absolute) |  | 01: 0; 02: 0.22; 03: 0.37; 04: 0.22; 06: 1.3; 09: 0.67; 10: 0.19; 11: 2.59; 13: 2.15; 14: 5.76; 15: 1.66
| Analytical results for Ni in OREAS 554 (Certified Value 33.7 ppm) | [22 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 1
1
| 1 | 1 | 37 | < 50 | < 100 | < 50 | 40 | 49.7 | 50 | 30 | < 50 | < 50 | NA | 33.8799 | 37 | 26 | 50 | < 50 | 31 | 31 |  |  |  | 1
| 1 | 2 | 26 | 100 | < 100 | 60 | 50 | 47.3 | 30 | 20 | < 50 | < 50 | NA | 33.7042 | 37 | 30 | 50 | < 50 | 29 | 28 |  |  |  | 0
| 1 | 3 | 30 | < 50 | < 100 | 60 | 60 | 47.2 | 30 | < 20 | < 50 | < 50 | NA | 37.5737 | 36 | 28 | 50 | < 50 | 28 | 44 |  |  |  | 16
| 1 | 4 | 28 | < 50 | < 100 | < 50 | 50 | 48.7 | 40 | 30 | < 50 | 60 | NA | 37.0205 | 39 | 30 | 50 | < 50 | 27 | 27 |  |  |  | 33.6704
| 1 | 5 | 26 | < 50 | < 100 | < 50 | 40 | 47.4 | 30 | < 20 | 50 | < 50 | NA | 35.4453 | 40 | 33 | 50 | < 50 | 31 | 18 |  |  |  | 59
| 1 | 6 | 27 | < 50 | < 100 | 50 | 60 | 50.1 | 20 | < 20 | < 50 | < 50 | NA | 35.7328 | 41 | 28 | 50 | < 50 | 28 | 19
| Mean |  | 29 | 100 |  | 56.6667 | 50 | 48.4 | 33.3333 | 26.6667 | 50 | 60 |  | 35.5594 | 38.3333 | 29.1667 | 50 |  | 29 | 27.8333
| Median |  | 27.5 | 100 |  | 60 | 50 | 48.05 | 30 | 30 | 50 | 60 |  | 35.5891 | 38 | 29 | 50 |  | 28.5 | 27.5
| Std Dev. |  | 01: 4.1952; 04: 5.7735; 06: 8.9443; 09: 1.29; 10: 10.328; 11: 5.7735; 16: 1.581; 18: 1.9664; 19: 2.4014; 20: 0; 23: 1.6733; 24: 9.4534
| Rel.Std.Dev. |  | 01: 14.47%; 04: 10.19%; 06: 17.89%; 09: 2.665%; 10: 30.98%; 11: 21.65%; 16: 4.446%; 18: 5.13%; 19: 8.233%; 20: 0%; 23: 5.77%; 24: 33.96%
| PDM3 |  | -13.87% | 197% |  | 68.3% | 48.5% | 43.75% | -1.001% | -20.8% | 48.5% | 78.2% |  | 5.61% | 13.85% | -13.38% | 48.5% |  | -13.87% | -17.34%
| Z-Score (Absolute) |  | 0.47 | 0.72 | 2.46 | NA | NA | 2.24 | NA | NA | NA | NA | NA | 0.45 | 0.83 | 0.45 | NA | 1.03 | 0.47 | 0.63
Results from laboratories 9, 10, 13, 14, 15, 16 & 23 were removed due to their 10 ppm reading resolution.
| Analytical results for P in OREAS 554 (Certified Value 0.034 wt.%) | [22 empty] | Certified
P | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 01; 02: 02; 03: 06; 04: 18; 06: 20; 09: 21; 10: 29; 11: 30 | wt.%
|  |  | 01: PF*OES; 02: PF*OES; 03: PF*OES/MS; 04: PF*OES; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: BF*OES/MS | 3
3
| 1 | 1 | 01: 0.03; 02: 0.03; 03: 0.04; 04: NA; 06: < 0.1; 09: 0.0304; 10: 0.03; 11: < 0.004 | 1
| 1 | 2 | 01: 0.03; 02: 0.03; 03: 0.04; 04: NA; 06: < 0.1; 09: 0.0346; 10: 0.03; 11: < 0.004 | 0
| 1 | 3 | 01: 0.04; 02: 0.04; 03: 0.04; 04: NA; 06: < 0.1; 09: 0.0406; 10: 0.03; 11: < 0.004 | 16
| 1 | 4 | 01: 0.03; 02: 0.03; 03: 0.04; 04: NA; 06: < 0.1; 09: 0.0389; 10: 0.03; 11: < 0.004 | 0.033883
| 1 | 5 | 01: 0.03; 02: 0.04; 03: 0.04; 04: NA; 06: < 0.1; 09: 0.0276; 10: 0.03; 11: < 0.004 | 60
| 1 | 6 | 01: 0.03; 02: 0.03; 03: 0.04; 04: NA; 06: < 0.1; 09: 0.0344; 10: 0.03; 11: < 0.004
| Mean |  | 01: 0.031667; 02: 0.033333; 03: 0.04; 09: 0.034417; 10: 0.03
| Median |  | 01: 0.03; 02: 0.03; 03: 0.04; 09: 0.0345; 10: 0.03
| Std Dev. |  | 01: 0.0040825; 02: 0.005164; 03: 0; 09: 0.0049187; 10: 0
| Rel.Std.Dev. |  | 01: 12.89%; 02: 15.49%; 03: 0%; 09: 14.29%; 10: 0%
| PDM3 |  | 01: -6.542%; 02: -1.623%; 03: 18.05%; 09: 1.574%; 10: -11.46%
| Z-Score (Absolute) |  | 01: 0.34; 02: 0; 03: 1.35; 04: NA; 06: 3.37; 09: 0.22; 10: 0.67; 11: 6.3
| Analytical results for Pb in OREAS 554 (Certified Value 18.6 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES |  | 1
1
| 1 | 1 | 18 | 28 | < 20 | 20 | 20 | < 300 | 14.7 | 19.8 | < 100 | < 100 | < 100 | 100 | NA | 19.6259 | 18 | < 20 | 22 | < 50 | 16 | 17.6 |  | 1
| 1 | 2 | 17 | 24 | < 20 | 20 | 20 | < 300 | 14.1 | 19.7 | < 100 | < 100 | < 100 | < 100 | NA | 17.0869 | 20 | < 20 | 21 | < 50 | 16 | 15.9 |  | 0
| 1 | 3 | 17 | < 20 | < 20 | 20 | 20 | < 300 | 13 | 19.5 | < 100 | 100 | < 100 | 100 | NA | 15.819 | 16 | < 20 | 22 | < 50 | 15 | 18.1 |  | 16
| 1 | 4 | 16 | < 20 | < 20 | 20 | 20 | < 300 | 14.2 | 18.2 | < 100 | < 100 | < 100 | < 100 | NA | 16.9286 | 15 | < 20 | 22 | < 50 | 15 | 28.6 |  | 18.6149
| 1 | 5 | 17 | < 20 | < 20 | 20 | 20 | < 300 | 14.9 | 18.2 | < 100 | < 100 | < 100 | 100 | NA | 15.8396 | 16 | < 20 | 24 | < 50 | 15 | 23.1 |  | 61
| 1 | 6 | 18 | 20 | < 20 | 20 | 20 | < 300 | 13.4 | 20.1 | < 100 | 100 | < 100 | 100 | NA | 15.4401 | 16 | < 20 | 22 | < 50 | 16 | 23.5
| Mean |  | 17.1667 | 24 |  | 20 | 20 |  | 14.05 | 19.25 |  | 100 |  | 100 |  | 16.79 | 16.8333 |  | 22.1667 |  | 15.5 | 21.1333
| Median |  | 17 | 24 |  | 20 | 20 |  | 14.15 | 19.6 |  | 100 |  | 100 |  | 16.3841 | 16 |  | 22 |  | 15.5 | 20.6
| Std Dev. |  | 0.75277 | 4 |  | 0 | 0 |  | 0.73417 | 0.83606 |  | 0 |  | 0 |  | 1.5374 | 1.8348 |  | 0.98319 |  | 0.54772 | 4.7811
| Rel.Std.Dev. |  | 4.385% | 16.67% |  | 0% | 0% |  | 5.225% | 4.343% |  | 0% |  | 0% |  | 9.157% | 10.9% |  | 4.435% |  | 3.534% | 22.62%
| PDM3 |  | -7.78% | 28.93% |  | 7.441% | 7.441% |  | -24.52% | 3.412% |  | 437.2% |  | 437.2% |  | -9.803% | -9.57% |  | 19.08% |  | -16.73% | 13.53%
| Z-Score (Absolute) |  | 0.42 | 0.45 | 1.5 | 0 | 0 | 19.48 | 0.89 | 0.11 | 4.5 | 6.99 | 4.5 | 9.49 | NA | 0.48 | 0.47 | 1.5 | 0.32 | 0.75 | 0.67 | 0.17
| Analytical results for Pr in OREAS 554 (Certified Value 7.86 ppm) | [22 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
3
| 1 | 1 | 01: 7.84; 02: 7.8; 03: 8; 04: 7.5; 06: 8.4; 09: 7.85; 10: 8.17; 11: 6.77; 13: 7.12; 14: 5.52; 15: 8.32 | 1
| 1 | 2 | 01: 7.7; 02: 8.2; 03: 8; 04: 7.5; 06: 8.1; 09: 7.72; 10: 8.89; 11: 6.73; 13: 7.18; 14: 5.61; 15: 8.06 | 0
| 1 | 3 | 01: 7.67; 02: 8; 03: 8; 04: 7.5; 06: 7.8; 09: 7.81; 10: 7.11; 11: 6.67; 13: 7.3; 14: 5.45; 15: 8.22 | 16
| 1 | 4 | 01: 7.51; 02: 8.2; 03: 8; 04: 7.5; 06: 9; 09: 7.73; 10: 7.59; 11: 6.73; 13: 7.1; 14: 5.36; 15: 9 | 7.8567
| 1 | 5 | 01: 7.65; 02: 7.8; 03: 8; 04: 8; 06: 8.4; 09: 7.69; 10: 7.27; 11: 7.22; 13: 7.22; 14: 5.63; 15: 8.85 | 62
| 1 | 6 | 01: 7.68; 02: 7.8; 03: 7.6; 04: 7.5; 06: 8.9; 09: 7.73; 10: 7.49; 11: 6.63; 13: 7.18; 14: 5.45; 15: 8.21
| Mean |  | 01: 7.675; 02: 7.9667; 03: 7.9333; 04: 7.5833; 06: 8.4333; 09: 7.755; 10: 7.7533; 11: 6.7917; 13: 7.1833; 14: 5.5033; 15: 8.4433
| Median |  | 01: 7.675; 02: 7.9; 03: 8; 04: 7.5; 06: 8.4; 09: 7.73; 10: 7.54; 11: 6.73; 13: 7.18; 14: 5.485; 15: 8.27
| Std Dev. |  | 01: 0.10559; 02: 0.19664; 03: 0.1633; 04: 0.20412; 06: 0.45898; 09: 0.061237; 10: 0.6647; 11: 0.21563; 13: 0.072019; 14: 0.10386; 15: 0.38516
| Rel.Std.Dev. |  | 01: 1.376%; 02: 2.468%; 03: 2.058%; 04: 2.692%; 06: 5.443%; 09: 0.7896%; 10: 8.573%; 11: 3.175%; 13: 1.003%; 14: 1.887%; 15: 4.562%
| PDM3 |  | 01: -2.312%; 02: 1.4%; 03: 0.9758%; 04: -3.479%; 06: 7.34%; 09: -1.294%; 10: -1.315%; 11: -13.56%; 13: -8.57%; 14: -29.95%; 15: 7.467%
| Z-Score (Absolute) |  | 01: 0.25; 02: 0.67; 03: 0.57; 04: 0.54; 06: 2.15; 09: 0.01; 10: 0; 11: 3.04; 13: 1.8; 14: 7.11; 15: 2.18
| Analytical results for Rb in OREAS 554 (Certified Value 89 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 28 | 29 | 30 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [6 empty] | 0
1
| 1 | 1 | 88 | 84.2 | 94.2 | 91.5 | 89 | 90.8 | 89.4 | NA | 101.2958 | 91 | 83 | 85.5 | 89 | 84.6 | 87.1 | [6 empty] | 1
| 1 | 2 | 88.9 | 84.9 | 94.2 | 92.5 | 90 | 94.5 | 88.6 | NA | 89.2572 | 90 | 86 | 85.5 | 86 | 83.1 | 92.6 | [6 empty] | 0
| 1 | 3 | 90.9 | 80.4 | 90.6 | 91.5 | 92 | 84.5 | 88.4 | NA | 88.1439 | 85.6 | 84 | 87.9 | 85 | 83.2 | 90.3 | [6 empty] | 16
| 1 | 4 | 89 | 81.1 | 91.2 | 92.5 | 90 | 89.1 | 88.3 | NA | 89.754 | 83 | 85 | 87.3 | 87 | 84.2 | 87.1 | [6 empty] | 88.8619
| 1 | 5 | 88.7 | 82.7 | 91.4 | 89 | 91 | 105 | 87.8 | NA | 91.2724 | 87.6 | 90 | 86.4 | 90 | 82.6 | 92.2 | [6 empty] | 63
| 1 | 6 | 88.9 | 83.1 | 91 | 89.5 | 86 | 88.5 | 87 | NA | 90.8888 | 89.6 | 85 | 90.3 | 88 | 85.3 | 89.4
| Mean |  | 89.0667 | 82.7333 | 92.1 | 91.0833 | 89.6667 | 92.0667 | 88.25 |  | 91.7687 | 87.8 | 85.5 | 87.15 | 87.5 | 83.8333 | 89.7833
| Median |  | 88.9 | 82.9 | 91.3 | 91.5 | 90 | 89.95 | 88.35 |  | 90.3214 | 88.6 | 85 | 86.85 | 87.5 | 83.7 | 89.85
| Std Dev. |  | 0.96885 | 1.7374 | 1.648 | 1.4972 | 2.0656 | 7.1214 | 0.80436 |  | 4.8019 | 3.0411 | 2.429 | 1.8174 | 1.8708 | 1.0328 | 2.3912
| Rel.Std.Dev. |  | 1.088% | 2.1% | 1.789% | 1.644% | 2.304% | 7.735% | 0.9115% |  | 5.233% | 3.464% | 2.841% | 2.085% | 2.138% | 1.232% | 2.663%
| PDM3 |  | 0.2304% | -6.897% | 3.644% | 2.5% | 0.9056% | 3.606% | -0.6886% |  | 3.271% | -1.195% | -3.783% | -1.927% | -1.533% | -5.659% | 1.037%
| Z-Score (Absolute) |  | 0.14 | 2.03 | 1.18 | 0.83 | 0.35 | 1.17 | 0.14 | NA | 1.07 | 0.29 | 1.08 | 0.52 | 0.4 | 1.65 | 0.39
| Analytical results for Re in OREAS 554 (Indicative Value < 0.1 ppm) | [22 empty] | Indicative
Re | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab | 1
| No | No | 01: 02; 02: 03; 03: 06; 04: 23 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*MS | 2
2
| 1 | 1 | 01: < 0.1; 02: < 1; 03: < 0.1; 04: < 0.1 | 1
| 1 | 2 | 01: < 0.1; 02: < 1; 03: < 0.1; 04: < 0.1 | 0
| 1 | 3 | 01: < 0.1; 02: < 1; 03: < 0.1; 04: < 0.1 | 16
| 1 | 4 | 01: < 0.1; 02: < 1; 03: < 0.1; 04: < 0.1 | < 0.1
| 1 | 5 | 01: < 0.1; 02: < 1; 03: < 0.1; 04: < 0.1 | 21
| 1 | 6 | 01: < 0.1; 02: < 1; 03: < 0.1; 04: < 0.1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 04
| Analytical results for S in OREAS 554 (Certified Value 0.769 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 23 | [8 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [8 empty] | 3
3
| 1 | 1 | 0.74 | 0.81 | 0.77 | 0.77 | 0.78 | 0.77 | 0.784 | 0.77 | 0.75 | 0.75 | 0.8 | NA | 0.85 | [8 empty] | 1
| 1 | 2 | 0.74 | 0.8 | 0.76 | 0.76 | 0.72 | 0.76 | 0.762 | 0.76 | 0.78 | 0.75 | 0.75 | NA | 0.86 | [8 empty] | 0
| 1 | 3 | 0.74 | 0.8 | 0.78 | 0.77 | 0.72 | 0.76 | 0.781 | 0.77 | 0.76 | 0.74 | 0.76 | NA | 0.86 | [8 empty] | 16
| 1 | 4 | 0.73 | 0.81 | 0.76 | 0.77 | 0.75 | 0.75 | 0.781 | 0.78 | 0.79 | 0.75 | 0.77 | NA | 0.86 | [8 empty] | 0.76948
| 1 | 5 | 0.71 | 0.81 | 0.77 | 0.76 | 0.72 | 0.75 | 0.784 | 0.78 | 0.78 | 0.77 | 0.78 | NA | 0.86 | [8 empty] | 64
| 1 | 6 | 0.75 | 0.8 | 0.78 | 0.77 | 0.71 | 0.76 | 0.767 | 0.78 | 0.82 | 0.77 | 0.76 | NA | 0.86
| Mean |  | 01: 0.735; 02: 0.805; 03: 0.77; 04: 0.76667; 06: 0.73333; 09: 0.75833; 10: 0.7765; 11: 0.77333; 13: 0.78; 14: 0.755; 15: 0.77; 18: 0.85833
| Median |  | 01: 0.74; 02: 0.805; 03: 0.77; 04: 0.77; 06: 0.72; 09: 0.76; 10: 0.781; 11: 0.775; 13: 0.78; 14: 0.75; 15: 0.765; 18: 0.86
| Std Dev. |  | 01: 0.013784; 02: 0.0054772; 03: 0.0089443; 04: 0.005164; 06: 0.026583; 09: 0.0075277; 10: 0.0095237; 11: 0.008165; 13: 0.024495; 14: 0.012247; 15: 0.017889; 18: 0.0040825
| Rel.Std.Dev. |  | 01: 1.875%; 02: 0.6804%; 03: 1.162%; 04: 0.6736%; 06: 3.625%; 09: 0.9927%; 10: 1.226%; 11: 1.056%; 13: 3.14%; 14: 1.622%; 15: 2.323%; 18: 0.4756%
| PDM3 |  | 01: -4.481%; 02: 4.616%; 03: 0.06714%; 04: -0.366%; 06: -4.698%; 09: -1.449%; 10: 0.9119%; 11: 0.5003%; 13: 1.367%; 14: -1.882%; 15: 0.06714%; 18: 11.55%
| Z-Score (Absolute) |  | 2.18 | 2.18 | 0 | 0.21 | 2.28 | 0.73 | 0.4 | 0.21 | 0.62 | 0.93 | 0 | NA | 5.5
| Analytical results for Sb in OREAS 554 (Certified Value 2.24 ppm) | [22 empty] | Certified
Sb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 18; 11: 19; 13: 20; 14: 21; 15: 29; 16: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: PF*MS; 13: PF*OES/MS; 14: PF*MS; 15: PF*MS; 16: PF*OES | 2
2
| 1 | 1 | 01: 1.6; 02: < 5; 03: < 10; 04: 4; 06: < 2; 09: 2.7; 10: NA; 11: 1.8823; 13: 3; 14: 2.2; 15: 2.8; 16: < 1 | 1
| 1 | 2 | 01: 1.8; 02: < 5; 03: < 10; 04: 4; 06: < 2; 09: 2.6; 10: NA; 11: 1.7629; 13: 3; 14: 2.3; 15: 2.6; 16: < 1 | 0
| 1 | 3 | 01: 1.9; 02: < 5; 03: < 10; 04: 5; 06: 2; 09: 2.7; 10: NA; 11: 1.7328; 13: 3; 14: 2; 15: 2.5; 16: < 1 | 16
| 1 | 4 | 01: 1.8; 02: < 5; 03: < 10; 04: 4; 06: 2; 09: 2.6; 10: NA; 11: 1.7416; 13: 3; 14: 2.3; 15: 2.8; 16: < 1 | 2.2418
| 1 | 5 | 01: 1.9; 02: < 5; 03: < 10; 04: 5; 06: 2; 09: 2.7; 10: NA; 11: 1.8027; 13: 3; 14: 2.3; 15: 3; 16: < 1 | 65
| 1 | 6 | 01: 1.7; 02: < 5; 03: < 10; 04: 4; 06: < 2; 09: 2.8; 10: NA; 11: 1.8314; 13: 3; 14: 2.2; 15: 2.7; 16: < 1
| Mean |  | 01: 1.7833; 04: 4.3333; 06: 2; 09: 2.6833; 11: 1.7923; 13: 3; 14: 2.2167; 15: 2.7333
| Median |  | 01: 1.8; 04: 4; 06: 2; 09: 2.7; 11: 1.7828; 13: 3; 14: 2.25; 15: 2.75
| Std Dev. |  | 01: 0.1169; 04: 0.5164; 06: 0; 09: 0.075277; 11: 0.057802; 13: 0; 14: 0.1169; 15: 0.17512
| Rel.Std.Dev. |  | 01: 6.555%; 04: 11.92%; 06: 0%; 09: 2.805%; 11: 3.225%; 13: 0%; 14: 5.274%; 15: 6.407%
| PDM3 |  | 01: -20.45%; 04: 93.3%; 06: -10.79%; 09: 19.7%; 11: -20.05%; 13: 33.82%; 14: -1.12%; 15: 21.93%
| Z-Score (Absolute) |  | 01: 0.82; 02: 0.2; 03: 3.79; 04: NA; 06: NA; 09: 0.47; 10: NA; 11: 0.81; 13: NA; 14: 0.2; 15: 0.54; 16: 2.66
Results from laboratories 6, 10 & 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 554 (Certified Value 7.98 ppm) | [22 empty] | Certified
Sc | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 06: 06; 09: 11; 10: 18; 11: 20; 13: 21; 14: 23; 15: 29; 16: 30 | ppm
|  |  | 01: PF*OES; 02: PF*OES; 03: PF*OES; 04: PF*OES; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: PF*OES/MS; 13: PF*MS; 14: PF*OES; 15: PF*MS; 16: BF*OES/MS | 2
2
| 1 | 1 | 01: 8; 02: < 10; 03: < 10; 04: < 10; 06: < 50; 09: 8.5; 10: NA; 11: 9; 13: 7; 14: < 10; 15: 7; 16: 6 | 1
| 1 | 2 | 01: 8; 02: < 10; 03: < 10; 04: < 10; 06: < 50; 09: 8.6; 10: NA; 11: 9; 13: 6; 14: < 10; 15: 8; 16: 6 | 0
| 1 | 3 | 01: 8; 02: < 10; 03: < 10; 04: < 10; 06: < 50; 09: 8.3; 10: NA; 11: 9; 13: 7; 14: < 10; 15: 7; 16: 7 | 16
| 1 | 4 | 01: 8; 02: < 10; 03: 10; 04: < 10; 06: < 50; 09: 8.3; 10: NA; 11: 9; 13: 7; 14: < 10; 15: 8; 16: 6 | 7.9771
| 1 | 5 | 01: 8; 02: < 10; 03: < 10; 04: < 10; 06: < 50; 09: 8.3; 10: NA; 11: 9; 13: 8; 14: < 10; 15: 7; 16: 6 | 66
| 1 | 6 | 01: 8; 02: < 10; 03: < 10; 04: < 10; 06: < 50; 09: 8.3; 10: NA; 11: 9; 13: 7; 14: < 10; 15: 7; 16: 6
| Mean |  | 01: 8; 03: 10; 09: 8.3833; 11: 9; 13: 7; 15: 7.3333; 16: 6.1667
| Median |  | 01: 8; 03: 10; 09: 8.3; 11: 9; 13: 7; 15: 7; 16: 6
| Std Dev. |  | 01: 0; 09: 0.13292; 11: 0; 13: 0.63246; 15: 0.5164; 16: 0.40825
| Rel.Std.Dev. |  | 01: 0%; 09: 1.585%; 11: 0%; 13: 9.035%; 15: 7.042%; 16: 6.62%
| PDM3 |  | 01: 0.2865%; 03: 25.36%; 09: 5.092%; 11: 12.82%; 13: -12.25%; 15: -8.071%; 16: -22.7%
| Z-Score (Absolute) |  | 01: 0.49; 02: 0.97; 03: 0.57; 04: 0.97; 06: 8.77; 09: 0.67; 10: NA; 11: 0.97; 13: 0; 14: 0.97; 15: 0.16; 16: 0.41
| Analytical results for Se in OREAS 554 (Indicative Value 13.3 ppm) | [22 empty] | Indicative
Se | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 02; 02: 03; 03: 06; 04: 10; 06: 18; 09: 19; 10: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 06: PF*OES; 09: PF*MS; 10: PF*OES | 1
1
| 1 | 1 | 01: < 20; 02: < 50; 03: 40; 04: 17; 06: NA; 09: 5.6517; 10: 2 | 1
| 1 | 2 | 01: < 20; 02: < 50; 03: 37; 04: 28; 06: NA; 09: 8.0389; 10: 2 | 0
| 1 | 3 | 01: < 20; 02: < 50; 03: 24; 04: 9; 06: NA; 09: 7.3681; 10: < 1 | 16
| 1 | 4 | 01: < 20; 02: < 50; 03: 34; 04: < 8; 06: NA; 09: 9.0775; 10: < 1 | 13.343
| 1 | 5 | 01: < 20; 02: < 50; 03: 35; 04: 11; 06: NA; 09: 6.251; 10: 8 | 22
| 1 | 6 | 01: < 20; 02: < 50; 03: 28; 04: < 8; 06: NA; 09: 6.2442; 10: < 1
| Mean |  | 03: 33; 04: 16.25; 09: 7.1052; 10: 4
| Median |  | 03: 34.5; 04: 14; 09: 6.8095; 10: 2
| Std Dev. |  | 03: 5.933; 04: 8.5391; 09: 1.2973; 10: 3.4641
| Rel.Std.Dev. |  | 03: 17.98%; 04: 52.55%; 09: 18.26%; 10: 86.6%
| PDM3 |  | 03: 147.3%; 04: 21.79%; 09: -46.75%; 10: -70.02%
| Z-Score (Absolute) |  | 01: 0.11; 02: 1.46; 03: 2.31; 04: 0.11; 06: NA; 09: 0.42; 10: 0.93
| Analytical results for Si in OREAS 554 (Certified Value 28.75 wt.%) | [22 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 13 | 14 | 15 | 16 | 19 | 20 | 23 | 24 | 30 | [6 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | [6 empty] | 2
3
| 1 | 1 | 25.2 | 28 | 29.1 | 29.2 | 28.84 | 28.9 | 28.09 | 29.45 | 28.09 | 29.31 | 27.9227 | 29.2 | 28.5 | > 15 | 28.8 | [6 empty] | 1
| 1 | 2 | 24.8 | 28.1 | 28.7 | 29.4 | 28.59 | 29.2 | 27.91 | 30.06 | 28 | 28.52 | 27.587 | 29.1 | 29.6 | > 15 | 28.84 | [6 empty] | 0
| 1 | 3 | 25.1 | 28.2 | 28.8 | 29.3 | 28.76 | 29.2 | 28.09 | 29.22 | 28.98 | 28.61 | 27.1783 | 29.5 | 29.1 | > 15 | 28.61 | [6 empty] | 16
| 1 | 4 | 24.4 | 27.8 | 28.6 | 29.5 | 29.12 | 29.1 | 28 | 29.12 | 28.98 | 28.8 | 27.4605 | 29.3 | 29 | > 15 | 28.33 | [6 empty] | 28.7481
| 1 | 5 | 24.6 | 28.1 | 29.3 | 29 | 29.07 | 29.3 | 28.09 | 30.15 | 28.61 | 28.61 | 27.3291 | 29.3 | 29.2 | > 15 | 28.61 | [6 empty] | 67
| 1 | 6 | 25 | 28.1 | 29.1 | 29.2 | 28.79 | 29.1 | 27.91 | 29.64 | 28.61 | 28.89 | 27.1978 | 30.2 | 29.4 | > 15 | 28.8
| Mean |  | 24.85 | 28.05 | 28.9333 | 29.2667 | 28.8617 | 29.1333 | 28.015 | 29.6067 | 28.545 | 28.79 | 27.4459 | 29.4333 | 29.1333 |  | 28.665
| Median |  | 24.9 | 28.1 | 28.95 | 29.25 | 28.815 | 29.15 | 28.045 | 29.545 | 28.61 | 28.705 | 27.3948 | 29.3 | 29.15 |  | 28.705
| Std Dev. |  | 0.30822 | 0.13784 | 0.27325 | 0.17512 | 0.19994 | 0.13663 | 0.088487 | 0.42725 | 0.42213 | 0.28921 | 0.28082 | 0.39833 | 0.37771 |  | 0.19254
| Rel.Std.Dev. |  | 1.24% | 0.4914% | 0.9444% | 0.5984% | 0.6928% | 0.469% | 0.3159% | 1.443% | 1.479% | 1.005% | 1.023% | 1.353% | 1.296% |  | 0.6717%
| PDM3 |  | -13.56% | -2.428% | 0.6444% | 1.804% | 0.3951% | 1.34% | -2.55% | 2.987% | -0.7064% | 0.1459% | -4.53% | 2.384% | 1.34% |  | -0.289%
| Z-Score (Absolute) |  | 7.14 | 1.39 | 0.2 | 0.79 | 0.07 | 0.55 | 1.45 | 1.4 | 0.5 | 0.07 | 2.48 | 1.09 | 0.55 | NA | 0.29
| Analytical results for Sm in OREAS 554 (Certified Value 4.93 ppm) | [22 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
3
| 1 | 1 | 01: 5; 02: 5; 03: 5.5; 04: 4.5; 06: 5.4; 09: 5.2; 10: 5; 11: 4.5; 13: 4.25; 14: 4; 15: 5.56 | 1
| 1 | 2 | 01: 4.9; 02: 5.5; 03: 5.5; 04: 5; 06: 4.5; 09: 5.1; 10: 6; 11: 4.6; 13: 4.3; 14: 3.9; 15: 5.75 | 0
| 1 | 3 | 01: 4.8; 02: 5; 03: 5.5; 04: 5; 06: 4.5; 09: 5.2; 10: 5; 11: 4.6; 13: 4.33; 14: 3.7; 15: 4.87 | 16
| 1 | 4 | 01: 4.8; 02: 5; 03: 5; 04: 5.5; 06: 5.5; 09: 5.2; 10: 5; 11: 4.5; 13: 4.19; 14: 3.8; 15: 5.33 | 4.9298
| 1 | 5 | 01: 5.1; 02: 5; 03: 5; 04: 5; 06: 4.9; 09: 5.2; 10: 5; 11: 4.9; 13: 4.23; 14: 3.7; 15: 5.2 | 68
| 1 | 6 | 01: 4.9; 02: 5; 03: 5; 04: 5; 06: 5.1; 09: 5.1; 10: 5; 11: 4.3; 13: 4.14; 14: 3.9; 15: 4.76
| Mean |  | 01: 4.9167; 02: 5.0833; 03: 5.25; 04: 5; 06: 4.9833; 09: 5.1667; 10: 5.1667; 11: 4.5667; 13: 4.24; 14: 3.8333; 15: 5.245
| Median |  | 01: 4.9; 02: 5; 03: 5.25; 04: 5; 06: 5; 09: 5.2; 10: 5; 11: 4.55; 13: 4.24; 14: 3.85; 15: 5.265
| Std Dev. |  | 01: 0.1169; 02: 0.20412; 03: 0.27386; 04: 0.31623; 06: 0.43089; 09: 0.05164; 10: 0.40825; 11: 0.19664; 13: 0.069857; 14: 0.12111; 15: 0.38454
| Rel.Std.Dev. |  | 01: 2.378%; 02: 4.016%; 03: 5.216%; 04: 6.325%; 06: 8.647%; 09: 0.9995%; 10: 7.902%; 11: 4.306%; 13: 1.648%; 14: 3.159%; 15: 7.332%
| PDM3 |  | 01: -0.2667%; 02: 3.114%; 03: 6.495%; 04: 1.424%; 06: 1.086%; 09: 4.804%; 10: 4.804%; 11: -7.366%; 13: -13.99%; 14: -22.24%; 15: 6.393%
| Z-Score (Absolute) |  | 01: 0.24; 02: 0.29; 03: 0.81; 04: 0.03; 06: 0.03; 09: 0.55; 10: NA; 11: 1.34; 13: 2.37; 14: 3.65; 15: 0.8
Results from laboratory 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 554 (Indicative Value 2.41 ppm) | [22 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 18 | 19 | 20 | 21 | 24 | 28 | 29 | 30 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | [6 empty] | 2
2
| 1 | 1 | 2 | < 100 | < 10 | < 10 | < 1 | < 50 | 4.7 | NA | 20.6302 | < 5 | < 5 | < 100 | < 5 | 1 | 2.7 | [6 empty] | 1
| 1 | 2 | 2 | < 100 | < 10 | < 10 | 2 | < 50 | 2.6 | NA | 16.7603 | < 5 | < 5 | < 100 | 8 | 2 | 2.1 | [6 empty] | 0
| 1 | 3 | 2 | < 100 | < 10 | < 10 | 2 | < 50 | 4.1 | NA | 14.7302 | < 5 | < 5 | < 100 | < 5 | 1 | 2.4 | [6 empty] | 16
| 1 | 4 | 2 | < 100 | < 10 | < 10 | 2 | < 50 | 4.9 | NA | 10.6858 | < 5 | < 5 | < 100 | < 5 | 2 | 2.5 | [6 empty] | 2.4083
| 1 | 5 | 2 | 100 | < 10 | < 10 | 2 | < 50 | 3.7 | NA | 14.8241 | < 5 | < 5 | < 100 | < 5 | 2 | 2.6 | [6 empty] | 23
| 1 | 6 | 2 | < 100 | < 10 | < 10 | 3 | < 50 | 4.4 | NA | 17.438 | < 5 | < 5 | < 100 | < 5 | 1 | 2.5
| Mean |  | 01: 2; 02: 100; 06: 2.2; 10: 4.0667; 13: 15.8448; 18: 8; 19: 1.5; 20: 2.4667
| Median |  | 01: 2; 02: 100; 06: 2; 10: 4.25; 13: 15.7922; 18: 8; 19: 1.5; 20: 2.5
| Std Dev. |  | 01: 0; 06: 0.44721; 10: 0.83586; 13: 3.3219; 19: 0.54772; 20: 0.20656
| Rel.Std.Dev. |  | 01: 0%; 06: 20.33%; 10: 20.55%; 13: 20.97%; 19: 36.51%; 20: 8.374%
| PDM3 |  | 01: -16.96%; 02: 4052%; 06: -8.651%; 10: 68.86%; 13: 557.9%; 18: 232.2%; 19: -37.72%; 20: 2.422%
| Z-Score (Absolute) |  | 0.78 | 24.41 | 0.56 | 0.56 | 0.82 | 9.5 | 0.15 | NA | 5.41 | 0.56 | 0.56 | 20.68 | 0.15 | 1 | 0.57
| Analytical results for Sr in OREAS 554 (Certified Value 71 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 | [5 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [5 empty] | 0
1
| 1 | 1 | 68 | 70 | 65 | 74 | < 100 | 71 | 80 | 71 | 87.037 | 69.5 | 73 | 90 | 70 | 74 | 60 | 75 | [5 empty] | 1
| 1 | 2 | 68 | 71 | 64 | 74 | < 100 | 84 | 81 | 71 | 80.8715 | 70.8 | 76 | 90 | 72 | 68 | 61 | 73 | [5 empty] | 0
| 1 | 3 | 69 | 73 | 67 | 72 | < 100 | 72 | 78 | 71 | 78.6308 | 71.7 | 72 | 90 | 73 | 70 | 58 | 78 | [5 empty] | 16
| 1 | 4 | 65 | 70 | 64 | 73 | < 100 | 72 | 80 | 71 | 80.3562 | 71.2 | 72 | 90 | 70 | 72 | 62 | 72 | [5 empty] | 71.0197
| 1 | 5 | 67 | 71 | 64 | 74 | < 100 | 76 | 80 | 70 | 82.8629 | 70.7 | 79 | 90 | 71 | 71 | 60 | 69 | [5 empty] | 69
| 1 | 6 | 68 | 74 | 64 | 71 | < 100 | 76 | 81 | 69 | 82.7873 | 73.8 | 71 | 90 | 72 | 72 | 58 | 71
| Mean |  | 67.5 | 71.5 | 64.6667 | 73 |  | 75.1667 | 80 | 70.5 | 82.091 | 71.2833 | 73.8333 | 90 | 71.3333 | 71.1667 | 59.8333 | 73
| Median |  | 68 | 71 | 64 | 73.5 |  | 74 | 80 | 71 | 81.8294 | 71 | 72.5 | 90 | 71.5 | 71.5 | 60 | 72.5
| Std Dev. |  | 1.3784 | 1.6432 | 1.2111 | 1.2649 |  | 4.8339 | 1.0954 | 0.83666 | 2.8987 | 1.4331 | 3.0605 | 0 | 1.2111 | 2.0412 | 1.6021 | 3.1623
| Rel.Std.Dev. |  | 2.042% | 2.298% | 1.873% | 1.733% |  | 6.431% | 1.369% | 1.187% | 3.531% | 2.01% | 4.145% | 0% | 1.698% | 2.868% | 2.678% | 4.332%
| PDM3 |  | -4.956% | 0.6763% | -8.945% | 2.788% |  | 5.839% | 12.64% | -0.7318% | 15.59% | 0.3712% | 3.962% | 26.73% | 0.4416% | 0.2069% | -15.75% | 2.788%
| Z-Score (Absolute) |  | 1.03 | 0.04 | 1.8 | 0.45 | 5.75 | 1.03 | 2.34 | 0.22 | 2.9 | 0.01 | 0.67 | NA | 0 | 0.04 | 3.1 | 0.45
Results from laboratory 23 were removed due to their 10 ppm reading resolution.
| Analytical results for Ta in OREAS 554 (Certified Value 0.7 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES | [7 empty] | 2
2
| 1 | 1 | 0.6 | 0.9 | < 0.5 | 0.5 | 4 | 0.9 | 2.1 | NA | 2.0366 | 1 | < 10 | < 3 | 0.7 | 0.67 | [7 empty] | 1
| 1 | 2 | 0.5 | 0.9 | < 0.5 | 0.5 | 5 | 0.8 | 2 | NA | 1.834 | 1 | < 10 | < 3 | 0.7 | 0.63 | [7 empty] | 0
| 1 | 3 | 0.6 | 1 | < 0.5 | 0.5 | 3 | 0.8 | 2.2 | NA | 4.6073 | 1 | < 10 | < 3 | 0.7 | 0.66 | [7 empty] | 16
| 1 | 4 | 0.5 | 0.8 | < 0.5 | 0.5 | 4 | 0.9 | 2 | NA | 2.9488 | 1 | < 10 | < 3 | 0.7 | 0.7 | [7 empty] | 0.69694
| 1 | 5 | 0.6 | 1.3 | < 0.5 | 0.5 | 3 | 0.9 | 2.1 | NA | 4.4974 | 1 | < 10 | < 3 | 0.7 | 0.81 | [7 empty] | 70
| 1 | 6 | 0.5 | 1 | < 0.5 | 0.5 | 3 | 0.7 | 2.1 | NA | 3.0593 | 1 | < 10 | < 3 | 0.7 | 0.6
| Mean |  | 01: 0.55; 02: 0.98333; 04: 0.5; 06: 3.6667; 09: 0.83333; 10: 2.0833; 13: 3.1639; 14: 1; 18: 0.7; 19: 0.67833
| Median |  | 01: 0.55; 02: 0.95; 04: 0.5; 06: 3.5; 09: 0.85; 10: 2.1; 13: 3.0041; 14: 1; 18: 0.7; 19: 0.665
| Std Dev. |  | 01: 0.054772; 02: 0.17224; 04: 0; 06: 0.8165; 09: 0.08165; 10: 0.075277; 13: 1.1797; 14: 0; 18: 1.2162e-16; 19: 0.073052
| Rel.Std.Dev. |  | 01: 9.959%; 02: 17.52%; 04: 0%; 06: 22.27%; 09: 9.798%; 10: 3.613%; 13: 37.29%; 14: 0%; 18: 1.737e-14%; 19: 10.77%
| PDM3 |  | 01: -21.08%; 02: 41.09%; 04: -28.26%; 06: 426.1%; 09: 19.57%; 10: 198.9%; 13: 354%; 14: 43.48%; 18: 0.4384%; 19: -2.67%
| Z-Score (Absolute) |  | 0.57 | 0.3 | 1.18 | 0.67 | NA | 0 | 2.53 | NA | 4.71 | NA | 8.43 | 1.35 | 0.27 | 0.31
Results from laboratories 6 & 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 554 (Certified Value 0.59 ppm) | [22 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
2
| 1 | 1 | 01: 0.61; 02: 0.6; 03: 0.8; 04: 0.7; 06: 0.6; 09: 0.79; 10: < 5; 11: 0.58; 13: 0.54; 14: 0.45; 15: 0.59 | 1
| 1 | 2 | 01: 0.61; 02: 0.8; 03: 0.6; 04: 0.8; 06: 0.6; 09: 0.77; 10: < 5; 11: 0.6; 13: 0.57; 14: 0.48; 15: 0.58 | 0
| 1 | 3 | 01: 0.62; 02: 0.6; 03: 0.6; 04: 0.7; 06: 0.6; 09: 0.76; 10: < 5; 11: 0.58; 13: 0.55; 14: 0.47; 15: 0.6 | 16
| 1 | 4 | 01: 0.61; 02: 0.6; 03: 0.6; 04: 0.6; 06: 0.6; 09: 0.76; 10: < 5; 11: 0.57; 13: 0.53; 14: 0.47; 15: 0.57 | 0.59194
| 1 | 5 | 01: 0.58; 02: 0.6; 03: 0.6; 04: 0.7; 06: 0.7; 09: 0.78; 10: < 5; 11: 0.65; 13: 0.52; 14: 0.47; 15: 0.62 | 71
| 1 | 6 | 01: 0.6; 02: 0.6; 03: 0.6; 04: 0.6; 06: 0.7; 09: 0.76; 10: < 5; 11: 0.56; 13: 0.54; 14: 0.45; 15: 0.69
| Mean |  | 01: 0.605; 02: 0.63333; 03: 0.63333; 04: 0.68333; 06: 0.63333; 09: 0.77; 11: 0.59; 13: 0.54167; 14: 0.465; 15: 0.60833
| Median |  | 01: 0.61; 02: 0.6; 03: 0.6; 04: 0.7; 06: 0.6; 09: 0.765; 11: 0.58; 13: 0.54; 14: 0.47; 15: 0.595
| Std Dev. |  | 01: 0.013784; 02: 0.08165; 03: 0.08165; 04: 0.075277; 06: 0.05164; 09: 0.012649; 11: 0.032249; 13: 0.017224; 14: 0.012247; 15: 0.043551
| Rel.Std.Dev. |  | 01: 2.278%; 02: 12.89%; 03: 12.89%; 04: 11.02%; 06: 8.154%; 09: 1.643%; 11: 5.466%; 13: 3.18%; 14: 2.634%; 15: 7.159%
| PDM3 |  | 01: 2.206%; 02: 6.992%; 03: 6.992%; 04: 15.44%; 06: 6.992%; 09: 30.08%; 11: -0.3285%; 13: -8.494%; 14: -21.45%; 15: 2.769%
| Z-Score (Absolute) |  | 01: 0; 02: NA; 03: NA; 04: NA; 06: NA; 09: 1.76; 10: 20.18; 11: 0.16; 13: 0.67; 14: 1.49; 15: 0.04
Results from laboratories 3, 4, 6 & 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 554 (Indicative Value < 1 ppm) | [22 empty] | Indicative
Te | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 02; 02: 03; 03: 06; 04: 10; 06: 18; 09: 19; 10: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES/MS; 06: PF*OES; 09: PF*MS; 10: PF*OES | 2
2
| 1 | 1 | 01: < 2; 02: < 1; 03: < 0.5; 04: < 6; 06: NA; 09: < 1; 10: < 0.1 | 1
| 1 | 2 | 01: < 2; 02: < 1; 03: < 0.5; 04: < 6; 06: NA; 09: < 1; 10: < 0.1 | 0
| 1 | 3 | 01: < 2; 02: < 1; 03: < 0.5; 04: < 6; 06: NA; 09: < 1; 10: < 0.1 | 16
| 1 | 4 | 01: < 2; 02: < 1; 03: < 0.5; 04: 7; 06: NA; 09: < 1; 10: < 0.1 | < 1
| 1 | 5 | 01: < 2; 02: < 1; 03: < 0.5; 04: < 6; 06: NA; 09: < 1; 10: < 0.1 | 19
| 1 | 6 | 01: < 2; 02: < 1; 03: < 0.5; 04: < 6; 06: NA; 09: < 1; 10: < 0.1
| Mean |  | 04: 7
| Median |  | 04: 7
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 01: 0.96; 02: 0; 03: 0.48; 04: 6.1; 06: NA; 09: 0; 10: 0.87
| Analytical results for Th in OREAS 554 (Certified Value 9.29 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 28 | 29 | 30 | [6 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [6 empty] | 2
3
| 1 | 1 | 9.7 | 8.6 | 9.5 | 9 | 8.5 | 10 | 10.1 | NA | 9.289 | 10 | 8.7 | 9.9 | 10 | 6.9 | 9.4 | [6 empty] | 1
| 1 | 2 | 8.8 | 8.8 | 9.5 | 9.5 | 9 | 9.6 | 9.9 | NA | 8.3085 | 10.3 | 9 | 9.8 | 9 | 6.3 | 9.83 | [6 empty] | 0
| 1 | 3 | 9.3 | 8.8 | 9.5 | 9.5 | 9 | 9.6 | 9.9 | NA | 7.3069 | 9.6 | 8.8 | 9.8 | 9 | 6.7 | 9.71 | [6 empty] | 16
| 1 | 4 | 8.9 | 8.7 | 9.5 | 9.5 | 9 | 10.2 | 9.8 | NA | 7.8526 | 8.9 | 8.9 | 10 | 10 | 6.5 | 10.9 | [6 empty] | 9.2941
| 1 | 5 | 9.2 | 8.8 | 9 | 9.5 | 8.5 | 9.9 | 9.8 | NA | 7.5378 | 9.8 | 13.3 | 9.7 | 9 | 6.7 | 9.89 | [6 empty] | 72
| 1 | 6 | 9.4 | 9 | 9 | 9.5 | 8 | 9.5 | 10.3 | NA | 6.9645 | 10.3 | 8.4 | 9.6 | 9 | 6.6 | 9.23
| Mean |  | 9.2167 | 8.7833 | 9.3333 | 9.4167 | 8.6667 | 9.8 | 9.9667 |  | 7.8765 | 9.8167 | 9.5167 | 9.8 | 9.3333 | 6.6167 | 9.8267
| Median |  | 9.25 | 8.8 | 9.5 | 9.5 | 8.75 | 9.75 | 9.9 |  | 7.6952 | 9.9 | 8.85 | 9.8 | 9 | 6.65 | 9.77
| Std Dev. |  | 0.33116 | 0.13292 | 0.2582 | 0.20412 | 0.40825 | 0.27568 | 0.19664 |  | 0.83108 | 0.52694 | 1.8649 | 0.14142 | 0.5164 | 0.20412 | 0.58442
| Rel.Std.Dev. |  | 3.593% | 1.513% | 2.766% | 2.168% | 4.711% | 2.813% | 1.973% |  | 10.55% | 5.368% | 19.6% | 1.443% | 5.533% | 3.085% | 5.947%
| PDM3 |  | -0.8329% | -5.495% | 0.4224% | 1.319% | -6.751% | 5.443% | 7.237% |  | -15.25% | 5.623% | 2.395% | 5.443% | 0.4224% | -28.81% | 5.73%
| Z-Score (Absolute) |  | 0.34 | 1.07 | 0.14 | 0 | 1.26 | 0.65 | 0.93 | NA | 2.6 | 0.67 | 0.17 | 0.65 | NA | 4.72 | 0.69
Results from laboratory 28 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 554 (Certified Value 0.228 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | BF*OES/MS |  | 3
3
| 1 | 1 | 0.22 | 0.24 | 0.24 | 0.23 | 0.24 | 0.24 | 0.23 | 0.23 | 0.222 | 0.216 | 0.222 | 0.24 | 0.2152 | 0.2296 | 0.22 | 0.2328 | 0.22 | 0.24 | 0.21 | 0.24 |  | 1
| 1 | 2 | 0.22 | 0.24 | 0.23 | 0.23 | 0.24 | 0.23 | 0.23 | 0.24 | 0.222 | 0.222 | 0.222 | 0.228 | 0.2109 | 0.2266 | 0.22 | 0.2273 | 0.23 | 0.23 | 0.21 | 0.246 |  | 0
| 1 | 3 | 0.22 | 0.25 | 0.22 | 0.24 | 0.24 | 0.23 | 0.23 | 0.23 | 0.222 | 0.216 | 0.228 | 0.234 | 0.2149 | 0.2242 | 0.22 | 0.2271 | 0.22 | 0.24 | 0.21 | 0.246 |  | 16
| 1 | 4 | 0.2 | 0.24 | 0.24 | 0.24 | 0.24 | 0.23 | 0.23 | 0.23 | 0.222 | 0.21 | 0.228 | 0.234 | 0.2149 | 0.226 | 0.22 | 0.2271 | 0.23 | 0.23 | 0.21 | 0.24 |  | 0.22793
| 1 | 5 | 0.22 | 0.28 | 0.24 | 0.24 | 0.24 | 0.22 | 0.22 | 0.23 | 0.222 | 0.222 | 0.228 | 0.228 | 0.2131 | 0.2242 | 0.22 | 0.2618 | 0.23 | 0.23 | 0.21 | 0.24 |  | 73
| 1 | 6 | 0.22 | 0.25 | 0.24 | 0.24 | 0.24 | 0.22 | 0.23 | 0.24 | 0.222 | 0.216 | 0.228 | 0.234 | 0.2079 | 0.2242 | 0.23 | 0.2238 | 0.23 | 0.23 | 0.2 | 0.246
| Mean |  | 0.21667 | 0.25 | 0.235 | 0.23667 | 0.24 | 0.22833 | 0.22833 | 0.23333 | 0.222 | 0.217 | 0.226 | 0.233 | 0.21282 | 0.2258 | 0.22167 | 0.23332 | 0.22667 | 0.23333 | 0.20833 | 0.243
| Median |  | 0.22 | 0.245 | 0.24 | 0.24 | 0.24 | 0.23 | 0.23 | 0.23 | 0.222 | 0.216 | 0.228 | 0.234 | 0.214 | 0.2251 | 0.22 | 0.2272 | 0.23 | 0.23 | 0.21 | 0.243
| Std Dev. |  | 0.008165 | 0.015492 | 0.0083666 | 0.005164 | 0 | 0.0075277 | 0.0040825 | 0.005164 | 0 | 0.0045166 | 0.0030984 | 0.0045166 | 0.0029082 | 0.0021354 | 0.0040825 | 0.014252 | 0.005164 | 0.005164 | 0.0040825 | 0.0032863
| Rel.Std.Dev. |  | 3.768% | 6.197% | 3.56% | 2.182% | 0% | 3.297% | 1.788% | 2.213% | 0% | 2.081% | 1.371% | 1.938% | 1.367% | 0.9457% | 1.842% | 6.109% | 2.278% | 2.213% | 1.96% | 1.352%
| PDM3 |  | -4.941% | 9.683% | 3.102% | 3.833% | 5.296% | 0.1771% | 0.1771% | 2.371% | -2.602% | -4.795% | -0.8467% | 2.224% | -6.631% | -0.9344% | -2.748% | 2.363% | -0.5542% | 2.371% | -8.598% | 6.612%
| Z-Score (Absolute) |  | 1.19 | 2.22 | 0.68 | 0.85 | 1.19 | 0 | 0 | 0.51 | 0.67 | 1.18 | 0.26 | 0.46 | 1.59 | 0.26 | 0.68 | 0.51 | 0.17 | 0.51 | 2.04 | 1.48
| Analytical results for Tl in OREAS 554 (Indicative Value 0.48 ppm) | [22 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 29 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*MS | [8 empty] | 2
2
| 1 | 1 | 0.6 | 0.6 | < 1 | < 2 | 1.5 | 0.6 | 0.5 | NA | < 1 | < 5 | < 0.5 | 0.6 | < 0.5 | [8 empty] | 1
| 1 | 2 | 0.5 | 0.6 | < 1 | < 2 | 1.4 | 0.3 | 0.5 | NA | < 1 | < 5 | 0.6 | 0.6 | < 0.5 | [8 empty] | 0
| 1 | 3 | 0.5 | 0.5 | < 1 | < 2 | 1.4 | 0.7 | 0.5 | NA | < 1 | < 5 | 0.5 | 0.5 | 0.5 | [8 empty] | 16
| 1 | 4 | 0.5 | < 0.5 | < 1 | < 2 | 1.3 | 0.7 | < 0.5 | NA | < 1 | < 5 | < 0.5 | 0.6 | < 0.5 | [8 empty] | 0.48429
| 1 | 5 | 0.6 | < 0.5 | < 1 | < 2 | 1.2 | 0.7 | < 0.5 | NA | < 1 | < 5 | 0.7 | 0.6 | < 0.5 | [8 empty] | 20
| 1 | 6 | < 0.5 | 0.7 | < 1 | < 2 | 1.1 | 0.5 | 0.5 | NA | < 1 | < 5 | 0.5 | 0.5 | 0.5
| Mean |  | 01: 0.54; 02: 0.6; 06: 1.3167; 09: 0.58333; 10: 0.5; 15: 0.575; 16: 0.56667; 18: 0.5
| Median |  | 01: 0.5; 02: 0.6; 06: 1.35; 09: 0.65; 10: 0.5; 15: 0.55; 16: 0.6; 18: 0.5
| Std Dev. |  | 01: 0.054772; 02: 0.08165; 06: 0.1472; 09: 0.16021; 10: 0; 15: 0.095743; 16: 0.05164; 18: 0
| Rel.Std.Dev. |  | 01: 10.14%; 02: 13.61%; 06: 11.18%; 09: 27.46%; 10: 0%; 15: 16.65%; 16: 9.113%; 18: 0%
| PDM3 |  | 01: 11.5%; 02: 23.89%; 06: 171.9%; 09: 20.45%; 10: 3.245%; 15: 18.73%; 16: 17.01%; 18: 3.245%
| Z-Score (Absolute) |  | 0.07 | 0.15 | 0 | 4.5 | 7.34 | 0.75 | 0.75 | NA | 0 | 17.98 | 0.3 | 0.6 | 1.5
| Analytical results for Tm in OREAS 554 (Certified Value 0.3 ppm) | [22 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
3
| 1 | 1 | 01: 0.3; 02: 0.2; 03: 0.4; 04: 0.4; 06: 0.3; 09: 0.32; 10: < 0.5; 11: 0.29; 13: 0.28; 14: 0.2; 15: 0.36 | 1
| 1 | 2 | 01: 0.29; 02: 0.2; 03: 0.2; 04: 0.4; 06: 0.4; 09: 0.33; 10: < 0.5; 11: 0.29; 13: 0.28; 14: 0.21; 15: 0.36 | 0
| 1 | 3 | 01: 0.29; 02: 0.4; 03: 0.2; 04: 0.4; 06: 0.3; 09: 0.31; 10: < 0.5; 11: 0.29; 13: 0.28; 14: 0.19; 15: 0.29 | 16
| 1 | 4 | 01: 0.29; 02: 0.2; 03: 0.4; 04: 0.3; 06: 0.3; 09: 0.32; 10: < 0.5; 11: 0.28; 13: 0.27; 14: 0.18; 15: 0.31 | 0.30033
| 1 | 5 | 01: 0.27; 02: 0.2; 03: 0.4; 04: 0.4; 06: 0.2; 09: 0.32; 10: < 0.5; 11: 0.3; 13: 0.26; 14: 0.2; 15: 0.34 | 74
| 1 | 6 | 01: 0.28; 02: 0.2; 03: 0.2; 04: 0.3; 06: 0.4; 09: 0.31; 10: < 0.5; 11: 0.29; 13: 0.29; 14: 0.21; 15: 0.32
| Mean |  | 01: 0.28667; 02: 0.23333; 03: 0.3; 04: 0.36667; 06: 0.31667; 09: 0.31833; 11: 0.29; 13: 0.27667; 14: 0.19833; 15: 0.33
| Median |  | 01: 0.29; 02: 0.2; 03: 0.3; 04: 0.4; 06: 0.3; 09: 0.32; 11: 0.29; 13: 0.28; 14: 0.2; 15: 0.33
| Std Dev. |  | 01: 0.010328; 02: 0.08165; 03: 0.10954; 04: 0.05164; 06: 0.075277; 09: 0.0075277; 11: 0.0063246; 13: 0.010328; 14: 0.01169; 15: 0.028284
| Rel.Std.Dev. |  | 01: 3.603%; 02: 34.99%; 03: 36.51%; 04: 14.08%; 06: 23.77%; 09: 2.365%; 11: 2.181%; 13: 3.733%; 14: 5.894%; 15: 8.571%
| PDM3 |  | 01: -4.55%; 02: -22.31%; 03: -0.111%; 04: 22.09%; 06: 5.438%; 09: 5.993%; 11: -3.441%; 13: -7.88%; 14: -33.96%; 15: 9.878%
| Z-Score (Absolute) |  | 01: 0; 02: NA; 03: NA; 04: NA; 06: NA; 09: 0.67; 10: 0.78; 11: 0.07; 13: 0.21; 14: 1.88; 15: 0.92
Results from laboratories 3, 4, 6 & 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 554 (Certified Value 14.3 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [7 empty] | 1
2
| 1 | 1 | 14.2 | 13.5 | 14.5 | 14 | 14.3 | 14.2 | NA | 12.3112 | 15.6 | 14.3 | 15.8 | 14 | 12.73 | 11 | [7 empty] | 1
| 1 | 2 | 14 | 14.5 | 14.5 | 14 | 14.6 | 14.1 | NA | 11.0621 | 16 | 14.07 | 16 | 13 | 13.75 | 9.8 | [7 empty] | 0
| 1 | 3 | 14.4 | 14.5 | 14.5 | 20 | 14.4 | 13.2 | NA | 10.4534 | 14.8 | 14.28 | 16.1 | 13 | 13.56 | 10.1 | [7 empty] | 16
| 1 | 4 | 13.8 | 14.5 | 14 | 14 | 14.6 | 13.4 | NA | 10.6984 | 15.2 | 14.3 | 16.6 | 13 | 13.77 | 8.8 | [7 empty] | 14.2717
| 1 | 5 | 14.5 | 14.5 | 15 | 14 | 13.7 | 13.7 | NA | 10.5859 | 15.1 | 16.26 | 15.8 | 13 | 13.73 | 8.99 | [7 empty] | 75
| 1 | 6 | 14.6 | 14 | 14 | 15 | 14.4 | 14.3 | NA | 10.1182 | 15.8 | 14.25 | 15.9 | 13 | 13.16 | 7.77
| Mean |  | 14.25 | 14.25 | 14.4167 | 15.1667 | 14.3333 | 13.8167 |  | 10.8715 | 15.4167 | 14.5767 | 16.0333 | 13.1667 | 13.45 | 9.41
| Median |  | 14.3 | 14.5 | 14.5 | 14 | 14.4 | 13.9 |  | 10.6422 | 15.4 | 14.29 | 15.95 | 13 | 13.645 | 9.395
| Std Dev. |  | 0.30822 | 0.41833 | 0.37639 | 2.4014 | 0.33267 | 0.4535 |  | 0.76977 | 0.45789 | 0.82923 | 0.30111 | 0.40825 | 0.42081 | 1.1309
| Rel.Std.Dev. |  | 2.163% | 2.936% | 2.611% | 15.83% | 2.321% | 3.282% |  | 7.081% | 2.97% | 5.689% | 1.878% | 3.101% | 3.129% | 12.02%
| PDM3 |  | -0.1518% | -0.1518% | 1.016% | 6.271% | 0.4321% | -3.188% |  | -23.82% | 8.023% | 2.137% | 12.34% | -7.743% | -5.757% | -34.07%
| Z-Score (Absolute) |  | 0 | 0 | 0.26 | NA | 0.13 | 0.67 | NA | 5.26 | 1.82 | 0.51 | 2.78 | NA | 1.24 | 7.53
Results from laboratories 6 & 28 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 554 (Certified Value 116 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | [7 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | [7 empty] | 0
0
| 1 | 1 | 117 | 116 | 150 | 150 | 100 | 121 | NA | 136.8295 | 108 | 109 | 130 | < 100 | 119 | 111 | [7 empty] | 1
| 1 | 2 | 119 | 116 | 150 | 150 | 100 | 120 | NA | 121.9236 | 114 | 108 | 130 | 113 | 113 | 110 | [7 empty] | 0
| 1 | 3 | 119 | 121 | 150 | 150 | 100 | 128 | NA | 115.1964 | 114 | 109 | 130 | 115 | 104 | 116 | [7 empty] | 16
| 1 | 4 | 115 | 114 | 150 | 150 | 100 | 122 | NA | 123.759 | 111 | 110 | 130 | 114 | 104 | 113 | [7 empty] | 115.9134
| 1 | 5 | 116 | 116 | 150 | 150 | 100 | 116 | NA | 120.3506 | 113 | 128 | 120 | 110 | 144 | 120 | [7 empty] | 76
| 1 | 6 | 117 | 120 | 150 | 150 | 100 | 125 | NA | 115.0628 | 113 | 108 | 130 | 111 | 139 | 117
| Mean |  | 117.1667 | 117.1667 | 150 | 150 | 100 | 122 |  | 122.187 | 112.1667 | 112 | 128.3333 | 112.6 | 120.5 | 114.5
| Median |  | 117 | 116 | 150 | 150 | 100 | 121.5 |  | 121.1371 | 113 | 109 | 130 | 113 | 116 | 114.5
| Std Dev. |  | 1.6021 | 2.7142 | 0 | 0 | 0 | 4.1473 |  | 7.9992 | 2.3166 | 7.874 | 4.0825 | 2.0736 | 17.3061 | 3.8341
| Rel.Std.Dev. |  | 1.367% | 2.316% | 0% | 0% | 0% | 3.399% |  | 6.547% | 2.065% | 7.03% | 3.181% | 1.842% | 14.36% | 3.349%
| PDM3 |  | 1.081% | 1.081% | 29.41% | 29.41% | -13.73% | 5.251% |  | 5.412% | -3.232% | -3.376% | 10.71% | -2.858% | 3.957% | -1.219%
| Z-Score (Absolute) |  | 0 | 0 | NA | NA | NA | 0.67 | NA | 0.7 | 0.7 | 0.72 | NA | 2.09 | 0.47 | 0.37
Results from laboratories 3, 4, 6 & 23 were removed due to their 10 ppm or greater reading resolution.
| Analytical results for W in OREAS 554 (Indicative Value 1.55 ppm) | [22 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 | [5 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | [5 empty] | 2
2
| 1 | 1 | 2 | 5 | < 5 | < 5 | < 10 | 2.2 | 1.9 | NA | 1.5321 | < 10 | < 100 | 7 | < 100 | < 5 | < 1 | 1 | [5 empty] | 1
| 1 | 2 | 2 | 6 | < 5 | < 5 | < 10 | 1.1 | 1.9 | NA | 1.2334 | < 10 | < 100 | 7 | < 100 | < 5 | 1 | < 1 | [5 empty] | 0
| 1 | 3 | 2 | 7 | < 5 | < 5 | < 10 | 1.4 | 1.8 | NA | 2.0852 | < 10 | < 100 | 7 | < 100 | < 5 | 1 | 2 | [5 empty] | 16
| 1 | 4 | 1 | 4 | < 5 | < 5 | < 10 | 2.5 | 1.7 | NA | 1.4326 | < 10 | < 100 | 7 | < 100 | < 5 | < 1 | < 1 | [5 empty] | 1.5512
| 1 | 5 | 1 | 9 | < 5 | < 5 | < 10 | 2.7 | 1.7 | NA | 1.6501 | < 10 | < 100 | 7 | < 100 | < 5 | 1 | < 1 | [5 empty] | 21
| 1 | 6 | 1 | 7 | < 5 | < 5 | < 10 | 1.4 | 1.8 | NA | 1.8091 | < 10 | < 100 | 7 | < 100 | < 5 | < 1 | < 1
| Mean |  | 01: 1.5; 02: 6.3333; 09: 1.8833; 10: 1.8; 13: 1.6237; 16: 7; 20: 1; 21: 1.5
| Median |  | 01: 1.5; 02: 6.5; 09: 1.8; 10: 1.8; 13: 1.5911; 16: 7; 20: 1; 21: 1.5
| Std Dev. |  | 01: 0.54772; 02: 1.7512; 09: 0.66758; 10: 0.089443; 13: 0.29844; 16: 0; 20: 0; 21: 0.70711
| Rel.Std.Dev. |  | 01: 36.51%; 02: 27.65%; 09: 35.45%; 10: 4.969%; 13: 18.38%; 16: 0%; 20: 0%; 21: 47.14%
| PDM3 |  | 01: -3.299%; 02: 308.3%; 09: 21.41%; 10: 16.04%; 13: 4.677%; 16: 351.3%; 20: -35.53%; 21: -3.299%
| Z-Score (Absolute) |  | 0.4 | 1.55 | 0 | 0 | 1.01 | 0.25 | 0.28 | NA | 0.35 | 1.01 | 19.22 | 1.82 | 19.22 | 0 | 0.71 | 0.67
| Analytical results for Y in OREAS 554 (Certified Value 20.1 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 10 | 11 | 18 | 20 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | PF*MS | PF*MS | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES | [7 empty] | 1
2
| 1 | 1 | 20.2 | 18.1 | 19 | 18 | 20.5 | 19.6 | 20.7 | 13.6 | 23 | 19.63 | 19.7 | 22 | 12.3 | 20.2 | [7 empty] | 1
| 1 | 2 | 20.1 | 18.6 | 19 | 20 | 18.5 | 19.9 | 20.7 | 13.3 | 23 | 19.82 | 21.5 | 21 | 12.3 | 20.2 | [7 empty] | 0
| 1 | 3 | 20.6 | 17.8 | 19 | 19 | 20.5 | 22.3 | 20.4 | 13.3 | 24 | 19.52 | 21.1 | 22 | 12.3 | 21 | [7 empty] | 16
| 1 | 4 | 19.8 | 18.8 | 20 | 20 | 20 | 21.4 | 20.1 | 13.3 | 20 | 19.54 | 19.3 | 21 | 13 | 20.4 | [7 empty] | 20.1249
| 1 | 5 | 19.1 | 19 | 20 | 19 | 19 | 19.3 | 20.2 | 13.6 | 21 | 19.49 | 20.8 | 22 | 13 | 21.3 | [7 empty] | 77
| 1 | 6 | 19.6 | 18.6 | 20 | 19 | 19.5 | 20.3 | 20 | 13.2 | 22 | 19.29 | 20.5 | 22 | 13 | 19.7
| Mean |  | 19.9 | 18.4833 | 19.5 | 19.1667 | 19.6667 | 20.4667 | 20.35 | 13.3833 | 22.1667 | 19.5483 | 20.4833 | 21.6667 | 12.65 | 20.4667
| Median |  | 19.95 | 18.6 | 19.5 | 19 | 19.75 | 20.1 | 20.3 | 13.3 | 22.5 | 19.53 | 20.65 | 22 | 12.65 | 20.3
| Std Dev. |  | 0.52154 | 0.44907 | 0.54772 | 0.75277 | 0.8165 | 1.157 | 0.30166 | 0.17224 | 1.472 | 0.17406 | 0.84004 | 0.5164 | 0.38341 | 0.58538
| Rel.Std.Dev. |  | 2.621% | 2.43% | 2.809% | 3.928% | 4.152% | 5.653% | 1.482% | 1.287% | 6.64% | 0.8904% | 4.101% | 2.383% | 3.031% | 2.86%
| PDM3 |  | -1.117% | -8.157% | -3.105% | -4.761% | -2.277% | 1.698% | 1.119% | -33.5% | 10.15% | -2.865% | 1.781% | 7.661% | -37.14% | 1.698%
| Z-Score (Absolute) |  | 0.12 | 1.28 | 0.28 | 0.61 | 0.12 | 0.67 | 0.56 | 6.32 | 2.35 | 0.23 | 0.69 | 1.86 | 7.04 | 0.67
| Analytical results for Yb in OREAS 554 (Certified Value 1.95 ppm) | [22 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 03; 03: 04; 04: 06; 06: 10; 09: 11; 10: 20; 11: 21; 13: 28; 14: 29; 15: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 06: PF*OES/MS; 09: PF*MS; 10: PF*OES/MS; 11: PF*MS; 13: PF*OES; 14: PF*MS; 15: PF*OES | 2
2
| 1 | 1 | 01: 2; 02: 2; 03: 2; 04: 2.2; 06: 1.4; 09: 2; 10: 2; 11: 2; 13: 1.86; 14: 1.3; 15: 2.07 | 1
| 1 | 2 | 01: 1.9; 02: 2; 03: 2; 04: 2; 06: 1.5; 09: 2.1; 10: 2; 11: 2; 13: 1.75; 14: 1.3; 15: 2.14 | 0
| 1 | 3 | 01: 2; 02: 2; 03: 2; 04: 2.8; 06: 1.9; 09: 2; 10: 2; 11: 1.9; 13: 1.83; 14: 1.4; 15: 2.34 | 16
| 1 | 4 | 01: 1.9; 02: 2.5; 03: 2; 04: 2; 06: 1.6; 09: 2; 10: 2; 11: 2; 13: 1.81; 14: 1.5; 15: 2 | 1.949
| 1 | 5 | 01: 1.9; 02: 2; 03: 2; 04: 2.1; 06: 1.4; 09: 2.1; 10: 2; 11: 1.9; 13: 1.69; 14: 1.4; 15: 2.24 | 78
| 1 | 6 | 01: 1.9; 02: 2; 03: 2; 04: 2.1; 06: 2.3; 09: 1.9; 10: 2; 11: 1.9; 13: 1.8; 14: 1.3; 15: 2.04
| Mean |  | 01: 1.9333; 02: 2.0833; 03: 2; 04: 2.2; 06: 1.6833; 09: 2.0167; 10: 2; 11: 1.95; 13: 1.79; 14: 1.3667; 15: 2.1383
| Median |  | 01: 1.9; 02: 2; 03: 2; 04: 2.1; 06: 1.55; 09: 2; 10: 2; 11: 1.95; 13: 1.805; 14: 1.35; 15: 2.105
| Std Dev. |  | 01: 0.05164; 02: 0.20412; 03: 0; 04: 0.30332; 06: 0.35449; 09: 0.075277; 10: 0; 11: 0.054772; 13: 0.060992; 14: 0.08165; 15: 0.12999
| Rel.Std.Dev. |  | 01: 2.671%; 02: 9.798%; 03: 0%; 04: 13.79%; 06: 21.06%; 09: 3.733%; 10: 0%; 11: 2.809%; 13: 3.407%; 14: 5.974%; 15: 6.079%
| PDM3 |  | 01: -0.8017%; 02: 6.895%; 03: 2.619%; 04: 12.88%; 06: -13.63%; 09: 3.474%; 10: 2.619%; 11: 0.05345%; 13: -8.156%; 14: -29.88%; 15: 9.717%
| Z-Score (Absolute) |  | 01: 0.07; 02: 0.56; 03: NA; 04: 1.05; 06: 1.12; 09: 0.28; 10: NA; 11: 0; 13: 0.67; 14: 2.46; 15: 0.79
Results from laboratories 4 & 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Zn in OREAS 554 (Certified Value 56 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 18 | 19 | 20 | 21 | 23 | 24 | 29 | 30 |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  | 0
0
| 1 | 1 | 61 | 69.06 | 50 | 60 | < 100 | < 100 | 80 | 35 | 40 | 30 | 200 | < 100 | NA | 54.033 | 55 | 76 | 100 | < 50 | 55 | 83 |  | 1
| 1 | 2 | 54 | 70.02 | 50 | 60 | < 100 | < 100 | 70 | 29 | 50 | 100 | 200 | < 100 | NA | 58.5255 | 52 | 68 | 100 | 57 | 55 | 74 |  | 0
| 1 | 3 | 56 | 60.46 | 50 | 70 | < 100 | < 100 | 70 | 29 | 50 | 40 | 200 | < 100 | NA | 45.0227 | 54 | < 50 | 100 | 51 | 56 | 79 |  | 16
| 1 | 4 | 52 | 49.96 | 100 | 60 | < 100 | < 100 | 60 | 32 | 50 | < 20 | 200 | < 100 | NA | 44.3139 | 59 | 87 | 100 | < 50 | 52 | 69 |  | 56.0093
| 1 | 5 | 56 | 55.93 | 50 | 60 | < 100 | < 100 | 90 | 31 | 40 | 40 | 200 | < 100 | NA | 43.8767 | 58 | < 50 | 100 | 54 | 54 | 84 |  | 79
| 1 | 6 | 59 | 56.44 | 50 | 60 | < 100 | < 100 | 80 | 31 | 50 | 150 | 200 | < 100 | NA | 59.5247 | 58 | 54 | 100 | 50 | 54 | 79
| Mean |  | 56.3333 | 60.3117 | 58.3333 | 61.6667 |  |  | 75 | 31.1667 | 46.6667 | 72 | 200 |  |  | 50.8827 | 56 | 71.25 | 100 | 53 | 54.3333 | 78
| Median |  | 56 | 58.45 | 50 | 60 |  |  | 75 | 31 | 50 | 40 | 200 |  |  | 49.5279 | 56.5 | 72 | 100 | 52.5 | 54.5 | 79
| Std Dev. |  | 3.266 | 7.9013 | 20.4124 | 4.0825 |  |  | 10.4881 | 2.2286 | 5.164 | 51.672 | 0 |  |  | 7.3429 | 2.7568 | 13.8894 | 0 | 3.1623 | 1.3663 | 5.6569
| Rel.Std.Dev. |  | 5.798% | 13.1% | 34.99% | 6.62% |  |  | 13.98% | 7.151% | 11.07% | 71.77% | 0% |  |  | 14.43% | 4.923% | 19.49% | 0% | 5.967% | 2.515% | 7.252%
| PDM3 |  | 0.5786% | 7.682% | 4.149% | 10.1% |  |  | 33.91% | -44.35% | -16.68% | 28.55% | 257.1% |  |  | -9.153% | -0.01653% | 27.21% | 78.54% | -5.373% | -2.992% | 39.26%
| Z-Score (Absolute) |  | 0.04 | 0.48 | 0.26 | 0.64 | 0.67 | 0.67 | 2.14 | 2.79 | 1.05 | 0.64 | 16.18 | 0.67 | NA | 0.58 | 0 | 0.02 | 4.94 | 1.39 | 0.19 | 2.47
| Analytical results for Zr in OREAS 554 (Certified Value 81 ppm) | [22 empty] | Certified
Zr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 06: Lab; 09: Lab | 1
| No | No | 01: 01; 02: 04; 03: 18; 04: 21; 06: 29; 09: 30 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES; 04: PF*MS; 06: PF*MS; 09: PF*OES | 0
0
| 1 | 1 | 01: 85.7; 02: 60; 03: NA; 04: 86.2; 06: 78; 09: 91 | 1
| 1 | 2 | 01: 84.7; 02: 70; 03: NA; 04: 83.8; 06: 76; 09: 99 | 0
| 1 | 3 | 01: 87.9; 02: 60; 03: NA; 04: 87.6; 06: 75; 09: 100 | 16
| 1 | 4 | 01: 82.5; 02: 70; 03: NA; 04: 101.5; 06: 74.8; 09: 98 | 81.1687
| 1 | 5 | 01: 87.1; 02: 70; 03: NA; 04: 83.7; 06: 76.3; 09: 91 | 80
| 1 | 6 | 01: 81.5; 02: 60; 03: NA; 04: 84.5; 06: 75.6; 09: 90
| Mean |  | 01: 84.9; 02: 65; 04: 87.8833; 06: 75.95; 09: 94.8333
| Median |  | 01: 85.2; 02: 65; 04: 85.35; 06: 75.8; 09: 94.5
| Std Dev. |  | 01: 2.5235; 02: 5.4772; 04: 6.8403; 06: 1.1554; 09: 4.6224
| Rel.Std.Dev. |  | 01: 2.972%; 02: 8.427%; 04: 7.783%; 06: 1.521%; 09: 4.874%
| PDM3 |  | 01: 4.597%; 02: -19.92%; 04: 8.272%; 06: -6.429%; 09: 16.83%
| Z-Score (Absolute) |  | 01: 0; 02: 1.5; 03: NA; 04: 0.22; 06: 0.67; 09: 0.75
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C621: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C639: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C730: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C748: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C766: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C876: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C933: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C951: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C970: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C988: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1006: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1044: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1063: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1081: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1099: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 554 (Certified Value 0.255 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.26 | 0.26 | < 0.5 | < 0.2 | 0.3 | < 1 | < 0.5 | 0.37 | 0.5 | 0.28 | 0.28 | 0.24 | 0.25 | 0.17 | 0.18 | 0.20354 | < 1 | < 0.02 | 0.3 | 0.25 | 0.26 | 0.23 | < 0.02 |  |  |  | 1
| 1 | 2 | 0.31 | 0.25 | < 0.5 | < 0.2 | 0.2 | < 1 | < 0.5 | 0.37 | 0.5 | 0.27 | 0.29 | 0.27 | 0.27 | 0.24 | 0.18 | 0.34889 | < 1 | < 0.02 | 0.3 | 0.24 | 0.28 | 0.25 | 0.07 |  |  |  | 0
| 1 | 3 | 0.24 | 0.26 | < 0.5 | < 0.2 | 0.2 | < 1 | < 0.5 | 0.39 | 0.6 | 0.26 | 0.28 | 0.26 | 0.28 | 0.21 | 0.22 | 0.21634 | < 1 | < 0.02 | 0.3 | 0.24 | 0.27 | 0.23 | 0.08 |  |  |  | 16
| 1 | 4 | 0.26 | 0.24 | < 0.5 | < 0.2 | 0.2 | < 1 | < 0.5 | 0.38 | 0.5 | 0.26 | 0.29 | 0.26 | 0.26 | 0.25 | 0.24 | 0.22965 | < 1 | < 0.02 | 0.3 | 0.23 | 0.26 | 0.25 | 0.1 |  |  |  | 0.25471
| 1 | 5 | 0.21 | 0.28 | < 0.5 | < 0.2 | 0.3 | < 1 | < 0.5 | 0.38 | 0.5 | 0.28 | 0.27 | 0.25 | 0.27 | 0.25 | 0.16 | 0.20815 | < 1 | < 0.02 | 0.3 | 0.23 | 0.26 | 0.23 | 0.05 |  |  |  | 82
| 1 | 6 | 0.29 | 0.31 | < 0.5 | < 0.2 | 0.3 | < 1 | < 0.5 | 0.36 | 0.5 | 0.27 | 0.3 | 0.24 | 0.26 | 0.24 | 0.17 | 0.23627 | < 1 | < 0.02 | 0.3 | 0.27 | 0.26 | 0.22 | 0.06
| Mean |  | 0.26167 | 0.26667 |  |  | 0.25 |  |  | 0.375 | 0.51667 | 0.27 | 0.285 | 0.25333 | 0.265 | 0.22667 | 0.19167 | 0.24047 |  |  | 0.3 | 0.24333 | 0.265 | 0.235 | 0.072
| Median |  | 0.26 | 0.26 |  |  | 0.25 |  |  | 0.375 | 0.5 | 0.27 | 0.285 | 0.255 | 0.265 | 0.24 | 0.18 | 0.22299 |  |  | 0.3 | 0.24 | 0.26 | 0.23 | 0.07
| Std Dev. |  | 0.035449 | 0.025033 |  |  | 0.054772 |  |  | 0.010488 | 0.040825 | 0.0089443 | 0.010488 | 0.012111 | 0.010488 | 0.031411 | 0.031252 | 0.054552 |  |  | 0 | 0.015055 | 0.0083666 | 0.012247 | 0.019235
| Rel.Std.Dev. |  | 13.55% | 9.387% |  |  | 21.91% |  |  | 2.797% | 7.902% | 3.313% | 3.68% | 4.781% | 3.958% | 13.86% | 16.31% | 22.69% |  |  | 0% | 6.187% | 3.157% | 5.212% | 26.72%
| PDM3 |  | 2.732% | 4.695% |  |  | -1.848% |  |  | 47.23% | 102.8% | 6.004% | 11.89% | -0.5397% | 4.041% | -11.01% | -24.75% | -5.589% |  |  | 17.78% | -4.466% | 4.041% | -7.738% | -71.73%
| Z-Score (Absolute) |  | 0.39 | 0.58 | 0.06 | 5.84 | NA | 9.57 | 0.06 | 4.75 | NA | 0.71 | 1.28 | 0.06 | 0.51 | 0.96 | 2.31 | 0.43 | 9.57 | 9.31 | NA | 0.32 | 0.51 | 0.64 | 7.32
Results from laboratories 5, 12 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 554 (Certified Value 4.32 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 4.35 | 4.2234 | 4.37 | 4.42 | 4.198 | 4.375 | 4.24 | 4.45 | 4.57 | 4.1 | 4.35 | 4.46 | 4.03 | 4.29 | 4.2 | 4.3235 | 4.3692 | 4.32 | 4.41 | 4.43 | 4.3065 | 4.29 | 4.15 | 4.377 |  |  | 1
| 1 | 2 | 4.37 | 4.2115 | 4.44 | 4.32 | 4.076 | 4.407 | 4.32 | 4.37 | 4.47 | 4.2 | 4.45 | 4.41 | 4.2 | 4.26 | 4.15 | 4.3182 | 4.3736 | 4.4 | 4.45 | 4.44 | 4.3263 | 4.33 | 4.15 | 4.361 |  |  | 0
| 1 | 3 | 4.39 | 4.2247 | 4.27 | 4.3 | 4.053 | 4.392 | 4.33 | 4.7 | 4.71 | 4.11 | 4.42 | 4.34 | 4.13 | 4.16 | 4.04 | 4.4158 | 4.112 | 4.4 | 4.55 | 4.31 | 4.3263 | 4.3 | 4.22 | 4.446 |  |  | 16
| 1 | 4 | 4.4 | 4.2473 | 4.25 | 4.37 | 4.135 | 4.355 | 4.25 | 4.47 | 4.5 | 4.2 | 4.32 | 4.37 | 4.09 | 4.25 | 3.97 | 4.2965 | 4.2944 | 4.28 | 4.45 | 4.34 | 4.4352 | 4.36 | 4.2 | 4.366 |  |  | 4.3167
| 1 | 5 | 4.37 | 4.23 | 4.35 | 4.34 | 4.163 | 4.348 | 4.27 | 4.54 | 4.77 | 4.16 | 4.21 | 4.4 | 4.15 | 4.2 | 4.15 | 4.4397 | 4.1866 | 4.39 | 4.54 | 4.51 | 4.2372 | 4.34 | 4.13 | 4.409 |  |  | 83
| 1 | 6 | 4.39 | 4.2524 | 4.36 | 4.34 | 4.175 | 4.416 | 4.27 | 4.65 | 4.49 | 4.09 | 4.24 | 4.29 | 4.2 | 4.36 | 4.02 | 4.297 | 4.213 | 4.38 | 4.42 | 4.65 | 4.4847 | 4.2 | 4.16 | 4.393
| Mean |  | 4.3783 | 4.2316 | 4.34 | 4.3483 | 4.1333 | 4.3822 | 4.28 | 4.53 | 4.585 | 4.1433 | 4.3317 | 4.3783 | 4.1333 | 4.2533 | 4.0883 | 4.3484 | 4.2581 | 4.3617 | 4.47 | 4.4467 | 4.3527 | 4.3033 | 4.1683 | 4.392
| Median |  | 4.38 | 4.2274 | 4.355 | 4.34 | 4.149 | 4.3835 | 4.27 | 4.505 | 4.535 | 4.135 | 4.335 | 4.385 | 4.14 | 4.255 | 4.095 | 4.3209 | 4.2537 | 4.385 | 4.45 | 4.435 | 4.3263 | 4.315 | 4.155 | 4.385
| Std Dev. |  | 0.018348 | 0.015495 | 0.069857 | 0.042151 | 0.057518 | 0.027622 | 0.036878 | 0.1257 | 0.12613 | 0.050067 | 0.095376 | 0.059133 | 0.065929 | 0.069761 | 0.090646 | 0.062843 | 0.10535 | 0.049967 | 0.060332 | 0.12307 | 0.090663 | 0.056804 | 0.034303 | 0.031837
| Rel.Std.Dev. |  | 0.4191% | 0.3662% | 1.61% | 0.9693% | 1.392% | 0.6303% | 0.8616% | 2.775% | 2.751% | 1.208% | 2.202% | 1.351% | 1.595% | 1.64% | 2.217% | 1.445% | 2.474% | 1.146% | 1.35% | 2.768% | 2.083% | 1.32% | 0.8229% | 0.7249%
| PDM3 |  | 1.427% | -1.973% | 0.5394% | 0.7325% | -4.248% | 1.516% | -0.8505% | 4.941% | 6.215% | -4.017% | 0.3464% | 1.427% | -4.248% | -1.468% | -5.291% | 0.7352% | -1.357% | 1.041% | 3.551% | 3.01% | 0.8336% | -0.31% | -3.437% | 1.744%
| Z-Score (Absolute) |  | 0.31 | 1.01 | 0.04 | 0.04 | 1.89 | 0.34 | 0.58 | 1.67 | 2.16 | 1.8 | 0.11 | 0.31 | 1.89 | 0.82 | 2.3 | 0.04 | 0.77 | 0.16 | 1.13 | 0.92 | 0.08 | 0.37 | 1.58 | 0.43
| Analytical results for As in OREAS 554 (Certified Value 66 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
1
| 1 | 1 | 66 | 68.5 | 69 | 62 | 64 | 66 | 71 | 60.9 | 67.8 | 67 | 64 | 68.5 | 67.8 | 68.4 | 66.2 | 63.6 | 125.7738 | 64 | 65 | 64 | 65.3 | 61.6 | 66 | 64.7 |  |  | 1
| 1 | 2 | 65 | 69.1 | 68 | 62 | 65 | 64 | 73 | 60.7 | 68 | 67 | 65.6 | 70 | 69.4 | 71.1 | 68.4 | 63.4 | 128.703 | 64 | 66 | 64 | 66.1 | 61.7 | 64 | 67.5 |  |  | 0
| 1 | 3 | 66 | 67 | 68 | 62 | 63 | 67 | 72 | 64.9 | 66.9 | 69 | 63.9 | 66.7 | 68.1 | 67.6 | 66.8 | 65.5 | 125.32 | 63 | 70 | 63 | 65.3 | 62.3 | 64 | 69.2 |  |  | 16
| 1 | 4 | 65 | 69.5 | 67 | 63 | 63 | 67 | 70 | 60.7 | 67.1 | 69 | 63.3 | 69.2 | 67.6 | 70 | 67.1 | 62.8 | 121.9498 | 62 | 68 | 65 | 67.6 | 61 | 64 | 69.6 |  |  | 65.9925
| 1 | 5 | 64 | 67.2 | 64 | 63 | 64 | 68 | 71 | 61.8 | 67.3 | 67 | 68.3 | 69.6 | 68.1 | 68.6 | 66.6 | 65.2 | 116.1513 | 62 | 70 | 63 | 66.3 | 62.6 | 64 | 66.5 |  |  | 84
| 1 | 6 | 66 | 69.5 | 66 | 59 | 65 | 64 | 69 | 60.7 | 67.1 | 69 | 68.5 | 66.4 | 68.6 | 71.2 | 64.3 | 63.6 | 118.4246 | 63 | 67 | 66 | 67.5 | 61.1 | 62 | 70.9
| Mean |  | 65.3333 | 68.4667 | 67 | 61.8333 | 64 | 66 | 71 | 61.6167 | 67.3667 | 68 | 65.6 | 68.4 | 68.2667 | 69.4833 | 66.5667 | 64.0167 | 122.7204 | 63 | 67.6667 | 64.1667 | 66.35 | 61.7167 | 64 | 68.0667
| Median |  | 65.5 | 68.8 | 67.5 | 62 | 64 | 66.5 | 71 | 60.8 | 67.2 | 68 | 64.8 | 68.85 | 68.1 | 69.3 | 66.7 | 63.6 | 123.6349 | 63 | 67.5 | 64 | 66.2 | 61.65 | 64 | 68.35
| Std Dev. |  | 0.8165 | 1.1219 | 1.7889 | 1.472 | 0.89443 | 1.6733 | 1.4142 | 1.6642 | 0.43665 | 1.0954 | 2.2996 | 1.5192 | 0.65013 | 1.5052 | 1.3397 | 1.0778 | 4.7765 | 0.89443 | 2.0656 | 1.169 | 1.0154 | 0.63692 | 1.2649 | 2.2704
| Rel.Std.Dev. |  | 1.25% | 1.639% | 2.67% | 2.381% | 1.398% | 2.535% | 1.992% | 2.701% | 0.6482% | 1.611% | 3.505% | 2.221% | 0.9523% | 2.166% | 2.012% | 1.684% | 3.892% | 1.42% | 3.053% | 1.822% | 1.53% | 1.032% | 1.976% | 3.336%
| PDM3 |  | -0.9988% | 3.749% | 1.527% | -6.302% | -3.019% | 0.01142% | 7.588% | -6.631% | 2.082% | 3.042% | -0.5947% | 3.648% | 3.446% | 5.29% | 0.8701% | -2.994% | 85.96% | -4.535% | 2.537% | -2.767% | 0.5418% | -6.479% | -3.019% | 3.143%
| Z-Score (Absolute) |  | 0.38 | 0.69 | 0.18 | 1.58 | 0.84 | 0.16 | 1.55 | 1.65 | 0.31 | 0.53 | 0.29 | 0.66 | 0.62 | 1.03 | 0.04 | 0.83 | 19.21 | 1.18 | 0.41 | 0.78 | 0.04 | 1.62 | 0.84 | 0.55
| Analytical results for B in OREAS 554 (Indicative Value 7.31 ppm) | [27 empty] | Indicative
B | Round | Replicate | 01: Lab; 02: Lab |  |  | 1
| No | No | 01: 12; 02: 19 |  |  | ppm
|  |  | 01: 4A*OES; 02: 4A*OES/MS |  |  | 2
2
| 1 | 1 | 01: 13; 02: 1.4777 |  |  | 1
| 1 | 2 | 01: 14; 02: 1.3173 |  |  | 0
| 1 | 3 | 01: 13; 02: 1.6951 |  |  | 16
| 1 | 4 | 01: 13; 02: 1.6748 |  |  | 7.3116
| 1 | 5 | 01: 13; 02: 1.4012 |  |  | 25
| 1 | 6 | 01: 13; 02: 1.1726
| Mean |  | 01: 13.1667; 02: 1.4564
| Median |  | 01: 13; 02: 1.4395
| Std Dev. |  | 01: 0.40825; 02: 0.20398
| Rel.Std.Dev. |  | 01: 3.101%; 02: 14.01%
| PDM3 |  | 01: 80.08%; 02: -80.08%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67
| Analytical results for Ba in OREAS 554 (Certified Value 1274 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
0
| 1 | 1 | 1212 | 1346.8 | 1280 | 1260 | 1277 | 1290 | 1253 | 298 | 1150 | 251 | 1300 | 1310 | 1100 | 1310 | 1140 | 1293 | 1544.2095 | 1258 | 1222.176 | 1380 | 857 | 1233 | 1265 | 1310 |  |  | 1
| 1 | 2 | 1235 | 1347.2 | 1290 | 1230 | 1267 | 1285 | 1268 | 320 | 1140 | 250 | 1350 | 1310 | 1380 | 1300 | 1220 | 1268 | 1576.7589 | 1283 | 1258.925 | 1320 | 812 | 1208 | 1252 | 1300 |  |  | 0
| 1 | 3 | 1238 | 1347.4 | 1310 | 1250 | 1258 | 1290 | 1263 | 213 | 1160 | 248 | 1330 | 1280 | 1300 | 1280 | 1050 | 1291 | 1589.3874 | 1286 | 1251.792 | 1310 | 818 | 1192 | 1281 | 1320 |  |  | 16
| 1 | 4 | 1239 | 1344 | 1290 | 1250 | 1278 | 1280 | 1241 | 162 | 1160 | 248 | 1310 | 1290 | 1340 | 1240 | 1150 | 1257 | 1549.3982 | 1266 | 1246.276 | 1330 | 884 | 1215 | 1273 | 1300 |  |  | 1274.3611
| 1 | 5 | 1232 | 1352.1 | 1270 | 1240 | 1257 | 1280 | 1244 | 279 | 1170 | 253 | 1260 | 1290 | 1350 | 1290 | 1200 | 1299 | 1524.9956 | 1291 | 1233.087 | 1290 | 814 | 1195 | 1315 | 1300 |  |  | 85
| 1 | 6 | 1208 | 1347.2 | 1280 | 1320 | 1272 | 1295 | 1241 | 243 | 1150 | 251 | 1240 | 1220 | 1370 | 1310 | 1120 | 1264 | 1557.6786 | 1251 | 1226.205 | 1350 | 693 | 1183 | 1302 | 1310
| Mean |  | 1227.3333 | 1347.45 | 1286.6667 | 1258.3333 | 1268.1667 | 1286.6667 | 1251.6667 | 252.5 | 1155 | 250.1667 | 1298.3333 | 1283.3333 | 1306.6667 | 1288.3333 | 1146.6667 | 1278.6667 | 1557.0714 | 1272.5 | 1239.7435 | 1330 | 813 | 1204.3333 | 1281.3333 | 1306.6667
| Median |  | 1233.5 | 1347.2 | 1285 | 1250 | 1269.5 | 1287.5 | 1248.5 | 261 | 1155 | 250.5 | 1305 | 1290 | 1345 | 1295 | 1145 | 1279.5 | 1553.5384 | 1274.5 | 1239.6815 | 1325 | 816 | 1201.5 | 1277 | 1305
| Std Dev. |  | 13.7064 | 2.6106 | 13.6626 | 31.8852 | 9.1524 | 6.0553 | 11.6905 | 58.5995 | 10.4881 | 1.9408 | 41.6733 | 33.2666 | 105.0079 | 26.3944 | 60.553 | 17.7163 | 23.1752 | 16.4286 | 14.7783 | 31.6228 | 65.4278 | 18.1292 | 23.4833 | 8.165
| Rel.Std.Dev. |  | 1.117% | 0.1937% | 1.062% | 2.534% | 0.7217% | 0.4706% | 0.934% | 23.21% | 0.9081% | 0.7758% | 3.21% | 2.592% | 8.036% | 2.049% | 5.281% | 1.386% | 1.488% | 1.291% | 1.192% | 2.378% | 8.048% | 1.505% | 1.833% | 0.6249%
| PDM3 |  | -3.69% | 5.735% | 0.9656% | -1.258% | -0.4861% | 0.9656% | -1.781% | -80.19% | -9.366% | -80.37% | 1.881% | 0.7041% | 2.535% | 1.096% | -10.02% | 0.3379% | 22.18% | -0.146% | -2.716% | 4.366% | -36.2% | -5.495% | 0.5471% | 2.535%
| Z-Score (Absolute) |  | 1.05 | 1.56 | 0.24 | 0.37 | 0.16 | 0.24 | 0.52 | 22.19 | 2.62 | 22.24 | 0.49 | 0.17 | 0.67 | 0.28 | 2.8 | 0.07 | 6.11 | 0.07 | 0.78 | 1.18 | 10.04 | 1.55 | 0.12 | 0.67
| Analytical results for Be in OREAS 554 (Certified Value 1.84 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 1.8 | 1.84 | 1.8 | 2 | 1.9 | 2 | < 5 | 1.9 | 2 | < 5 | 2.1 | 1.78 | 1.83 | 1.79 | 1.74 | 1.83 | 1.6889 | 1.8 | 1.7994 | 1.9 | 2 | 1.89 | 1.9 | 1.7 |  |  | 1
| 1 | 2 | 1.8 | 1.88 | 1.8 | 2 | 1.9 | 2 | < 5 | 1.9 | 2.1 | < 5 | 2.01 | 1.76 | 1.91 | 1.74 | 1.77 | 1.83 | 1.592 | 2 | 1.789 | 1.9 | 2 | 1.9 | 1.8 | 1.8 |  |  | 0
| 1 | 3 | 1.8 | 1.89 | 1.8 | 2 | 1.9 | 2 | < 5 | 2 | 2 | < 5 | 2.06 | 1.8 | 1.8 | 1.68 | 1.63 | 1.88 | 1.5977 | 2 | 1.8145 | 1.9 | 2 | 1.91 | 1.8 | 1.6 |  |  | 16
| 1 | 4 | 1.8 | 1.8 | 1.7 | 2 | 1.9 | 2 | < 5 | 1.9 | 2.2 | < 5 | 1.96 | 1.77 | 1.82 | 1.76 | 1.73 | 1.81 | 1.6213 | 2.1 | 1.7378 | 1.8 | 2 | 1.91 | 2 | 1.8 |  |  | 1.8399
| 1 | 5 | 1.8 | 1.79 | 1.7 | 2 | 1.9 | 2 | < 5 | 2.1 | 2.1 | < 5 | 1.92 | 1.72 | 1.78 | 1.72 | 1.63 | 1.87 | 1.6094 | 1.8 | 1.8156 | 1.8 | 2 | 1.97 | 2 | 1.7 |  |  | 86
| 1 | 6 | 1.8 | 1.79 | 1.8 | 2 | 1.9 | 2 | < 5 | 2 | 2.1 | < 5 | 1.85 | 1.87 | 1.77 | 1.69 | 1.71 | 1.83 | 1.6055 | 2 | 1.8162 | 1.9 | 2 | 1.92 | 1.9 | 1.7
| Mean |  | 1.8 | 1.8317 | 1.7667 | 2 | 1.9 | 2 |  | 1.9667 | 2.0833 |  | 1.9833 | 1.7833 | 1.8183 | 1.73 | 1.7017 | 1.8417 | 1.6191 | 1.95 | 1.7954 | 1.8667 | 2 | 1.9167 | 1.9 | 1.7167
| Median |  | 1.8 | 1.82 | 1.8 | 2 | 1.9 | 2 |  | 1.95 | 2.1 |  | 1.985 | 1.775 | 1.81 | 1.73 | 1.72 | 1.83 | 1.6074 | 2 | 1.807 | 1.9 | 2 | 1.91 | 1.9 | 1.7
| Std Dev. |  | 0 | 0.04535 | 0.05164 | 0 | 2.4324e-16 | 0 |  | 0.08165 | 0.075277 |  | 0.092232 | 0.050067 | 0.050365 | 0.041952 | 0.058793 | 0.027142 | 0.03564 | 0.12247 | 0.030275 | 0.05164 | 0 | 0.028048 | 0.089443 | 0.075277
| Rel.Std.Dev. |  | 0% | 2.476% | 2.923% | 0% | 1.28e-14% | 0% |  | 4.152% | 3.613% |  | 4.65% | 2.807% | 2.77% | 2.425% | 3.455% | 1.474% | 2.201% | 6.281% | 1.686% | 2.766% | 0% | 1.463% | 4.708% | 4.385%
| PDM3 |  | -2.169% | -0.4476% | -3.98% | 8.701% | 3.266% | 8.701% |  | 6.89% | 13.23% |  | 7.796% | -3.075% | -1.172% | -5.973% | -7.513% | 0.09592% | -12% | 5.984% | -2.417% | 1.455% | 8.701% | 4.172% | 3.266% | -6.698%
| Z-Score (Absolute) |  | 0.37 | 0.09 | 0.67 | NA | 0.52 | NA | 5.92 | 1.12 | 2.17 | 5.92 | 1.27 | 0.52 | 0.21 | 1 | 1.26 | 0 | 2 | 0.97 | 0.42 | 0.22 | NA | 0.67 | 0.52 | 1.12
Results from laboratories 4, 6 & 24 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 554 (Certified Value 2.74 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | < 5 | 2.8 | 2.8 | 2.5 | 2.84 | 3.2 | 2.9 | 2.83 | 3.58 | < 5 | 2.95 | 3.01 | 2.49 | 2.62 | 2.67 | 2.71 | 2.9827 | 2.62 | 2.7139 | 2.48 | 3.05 | 2.76 | < 5 | 2.1 |  |  | 1
| 1 | 2 | < 5 | 2.88 | 2.8 | 2.4 | 2.78 | 3 | 2.5 | 2.87 | 3.47 | < 5 | 3.01 | 2.94 | 2.63 | 2.58 | 2.84 | 2.67 | 3.0406 | 2.68 | 2.6072 | 2.59 | 2.82 | 2.75 | < 5 | 2.2 |  |  | 0
| 1 | 3 | < 5 | 2.89 | 2.7 | 2.4 | 2.79 | 3.1 | 2.5 | 3.03 | 3.5 | < 5 | 3.02 | 2.84 | 2.59 | 2.54 | 2.62 | 2.63 | 3.0712 | 2.76 | 2.6855 | 2.45 | 2.63 | 2.8 | < 5 | 2.1 |  |  | 16
| 1 | 4 | < 5 | 2.94 | 2.8 | 2.4 | 2.82 | 3.1 | 2.5 | 2.85 | 3.37 | < 5 | 2.99 | 2.86 | 2.55 | 2.6 | 2.7 | 2.61 | 2.9149 | 2.66 | 2.564 | 2.49 | 2.65 | 2.78 | < 5 | 2.1 |  |  | 2.7354
| 1 | 5 | < 5 | 2.84 | 2.8 | 2.5 | 2.85 | 3.2 | 2.5 | 2.9 | 3.26 | < 5 | 2.79 | 2.81 | 2.58 | 2.57 | 2.56 | 2.65 | 2.8511 | 2.67 | 2.6669 | 2.59 | 2.76 | 2.86 | < 5 | 2.4 |  |  | 87
| 1 | 6 | < 5 | 2.79 | 2.8 | 2.4 | 2.77 | 3 | 2.4 | 2.81 | 3.29 | < 5 | 2.85 | 2.78 | 2.6 | 2.62 | 2.64 | 2.65 | 2.9438 | 2.58 | 2.7501 | 2.69 | 2.59 | 2.77 | 5 | 2
| Mean |  |  | 2.8567 | 2.7833 | 2.4333 | 2.8083 | 3.1 | 2.55 | 2.8817 | 3.4117 |  | 2.935 | 2.8733 | 2.5733 | 2.5883 | 2.6717 | 2.6533 | 2.9674 | 2.6617 | 2.6646 | 2.5483 | 2.75 | 2.7867 | 5 | 2.15
| Median |  |  | 2.86 | 2.8 | 2.4 | 2.805 | 3.1 | 2.5 | 2.86 | 3.42 |  | 2.97 | 2.85 | 2.585 | 2.59 | 2.655 | 2.65 | 2.9632 | 2.665 | 2.6762 | 2.54 | 2.705 | 2.775 | 5 | 2.1
| Std Dev. |  |  | 0.057504 | 0.040825 | 0.05164 | 0.033116 | 0.089443 | 0.17607 | 0.079099 | 0.12576 |  | 0.094181 | 0.086178 | 0.048442 | 0.031252 | 0.095167 | 0.034448 | 0.0815 | 0.0608 | 0.068703 | 0.090866 | 0.17029 | 0.039833 |  | 0.13784
| Rel.Std.Dev. |  |  | 2.013% | 1.467% | 2.122% | 1.179% | 2.885% | 6.905% | 2.745% | 3.686% |  | 3.209% | 2.999% | 1.882% | 1.207% | 3.562% | 1.298% | 2.747% | 2.284% | 2.578% | 3.566% | 6.193% | 1.429% |  | 6.411%
| PDM3 |  |  | 4.433% | 1.752% | -11.04% | 2.666% | 13.33% | -6.778% | 5.347% | 24.72% |  | 7.297% | 5.042% | -5.925% | -5.376% | -2.33% | -3% | 8.48% | -2.696% | -2.589% | -6.839% | 0.5337% | 1.874% | 82.79% | -21.4%
| Z-Score (Absolute) |  | 0.88 | 0.61 | 0.3 | 1.16 | 0.41 | 1.62 | 0.67 | 0.71 | 2.92 | 0.88 | 0.93 | 0.68 | 0.57 | 0.51 | 0.16 | 0.24 | 1.07 | 0.21 | 0.19 | 0.68 | 0.16 | 0.32 | 0.86 | 2.34
| Analytical results for Ca in OREAS 554 (Certified Value 3.96 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 3.79 | 3.7679 | 4.01 | 3.94 | 4.121 | 4.0355 | 3.87 | 4.06 | 3.79 | 4.27 | 4.02 | 4.2 | 3.84 | 3.98 | 4.06 | 4.0809 | 4.1021 | 3.96 | 3.83 | 4.01 | 3.822 | 3.96 | 3.85 | 3.809 |  |  | 1
| 1 | 2 | 3.85 | 3.8123 | 3.98 | 3.92 | 3.997 | 4.018 | 3.92 | 3.84 | 3.8 | 4.22 | 4.11 | 4.13 | 4.01 | 3.95 | 4.01 | 4.0515 | 4.1216 | 4.08 | 3.88 | 3.9 | 3.822 | 3.94 | 3.96 | 3.852 |  |  | 0
| 1 | 3 | 3.83 | 3.7273 | 3.97 | 3.92 | 3.995 | 4.0265 | 3.93 | 4.29 | 3.79 | 4.23 | 4.1 | 4.02 | 3.93 | 3.9 | 3.93 | 4.1703 | 4.061 | 4.01 | 3.94 | 3.81 | 3.8612 | 3.95 | 3.9 | 3.924 |  |  | 16
| 1 | 4 | 3.86 | 3.8066 | 3.9 | 3.97 | 4.036 | 4.014 | 3.88 | 4.01 | 3.91 | 4.18 | 3.99 | 4.09 | 3.88 | 3.96 | 3.87 | 4.0667 | 4.0193 | 3.95 | 3.87 | 3.83 | 3.9886 | 4 | 3.95 | 3.859 |  |  | 3.96
| 1 | 5 | 3.84 | 3.8139 | 3.97 | 3.92 | 4.064 | 4.0305 | 3.87 | 4.18 | 3.86 | 4.12 | 3.87 | 4.08 | 3.92 | 3.9 | 4.02 | 4.1245 | 3.9272 | 3.92 | 3.94 | 4.06 | 3.7828 | 3.95 | 3.84 | 3.838 |  |  | 88
| 1 | 6 | 3.85 | 3.7871 | 3.99 | 3.94 | 4.105 | 4.0435 | 3.88 | 4.03 | 3.9 | 4.2 | 3.91 | 3.88 | 3.98 | 4.05 | 3.9 | 4.0822 | 3.9411 | 4 | 3.83 | 3.98 | 4.018 | 3.88 | 3.88 | 3.888
| Mean |  | 3.8367 | 3.7858 | 3.97 | 3.935 | 4.053 | 4.028 | 3.8917 | 4.0683 | 3.8417 | 4.2033 | 4 | 4.0667 | 3.9267 | 3.9567 | 3.965 | 4.096 | 4.0287 | 3.9867 | 3.8817 | 3.9317 | 3.8824 | 3.9467 | 3.8967 | 3.8617
| Median |  | 3.845 | 3.7969 | 3.975 | 3.93 | 4.05 | 4.0285 | 3.88 | 4.045 | 3.83 | 4.21 | 4.005 | 4.085 | 3.925 | 3.955 | 3.97 | 4.0816 | 4.0401 | 3.98 | 3.875 | 3.94 | 3.8416 | 3.95 | 3.89 | 3.8555
| Std Dev. |  | 0.025033 | 0.033671 | 0.037417 | 0.019748 | 0.053363 | 0.010964 | 0.026394 | 0.15407 | 0.055648 | 0.050859 | 0.09757 | 0.10912 | 0.062503 | 0.056095 | 0.075565 | 0.043795 | 0.081385 | 0.056451 | 0.049565 | 0.10108 | 0.097295 | 0.038816 | 0.050067 | 0.040033
| Rel.Std.Dev. |  | 0.6525% | 0.8894% | 0.9425% | 0.5019% | 1.317% | 0.2722% | 0.6782% | 3.787% | 1.449% | 1.21% | 2.439% | 2.683% | 1.592% | 1.418% | 1.906% | 1.069% | 2.02% | 1.416% | 1.277% | 2.571% | 2.506% | 0.9835% | 1.285% | 1.037%
| PDM3 |  | -3.115% | -4.398% | 0.2518% | -0.6321% | 2.348% | 1.716% | -1.726% | 2.735% | -2.989% | 6.144% | 1.009% | 2.693% | -0.8425% | -0.08494% | 0.1255% | 3.434% | 1.734% | 0.6726% | -1.979% | -0.7163% | -1.96% | -0.3375% | -1.6% | -2.484%
| Z-Score (Absolute) |  | 1.11 | 1.61 | 0.18 | 0.16 | 0.98 | 0.74 | 0.58 | 1.13 | 1.07 | 2.44 | 0.47 | 1.11 | 0.24 | 0.05 | 0.13 | 1.4 | 0.75 | 0.34 | 0.68 | 0.19 | 0.67 | 0.05 | 0.53 | 0.87
| Analytical results for Cd in OREAS 554 (Certified Value 0.21 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 0.23 | 0.26 | < 0.5 | < 0.5 | 0.19 | < 0.5 | < 0.5 | 0.2 | 0.15 | 0.5 | 0.22 | 0.19 | 0.2 | 0.23 | 0.2 | 0.22 | 0.23443 | 0.19 | 0.22 | 0.20856 | 0.2 | 0.17 | 0.21 | 0.18 | 0.2 |  | 1
| 1 | 2 | 0.2 | 0.22 | < 0.5 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.2 | 0.18 | 0.5 | 0.23 | 0.21 | 0.2 | 0.22 | 0.18 | 0.22 | 0.34174 | 0.19 | 0.26 | 0.17019 | 0.21 | 0.2 | 0.22 | 0.2 | 0.2 |  | 0
| 1 | 3 | 0.19 | 0.23 | < 0.5 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.2 | 0.15 | 0.5 | 0.26 | 0.18 | 0.22 | 0.21 | 0.18 | 0.2 | 0.23306 | 0.19 | 0.21 | 0.22017 | 0.23 | 0.21 | 0.22 | 0.19 | 0.2 |  | 16
| 1 | 4 | 0.24 | 0.2 | < 0.5 | < 0.5 | 0.21 | < 0.5 | < 0.5 | 0.2 | 0.13 | 0.6 | 0.19 | 0.21 | 0.19 | 0.19 | 0.2 | 0.19 | 0.24152 | 0.19 | 0.21 | 0.22199 | 0.22 | 0.19 | 0.23 | 0.18 | 0.3 |  | 0.20734
| 1 | 5 | 0.24 | 0.23 | < 0.5 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.2 | 0.15 | 0.5 | 0.45 | 0.22 | 0.21 | 0.2 | 0.19 | 0.21 | 0.22602 | 0.19 | 0.18 | 0.1552 | 0.22 | 0.17 | 0.22 | 0.2 | 0.3 |  | 89
| 1 | 6 | 0.18 | 0.24 | < 0.5 | < 0.5 | 0.19 | < 0.5 | < 0.5 | 0.2 | 0.13 | 0.6 | 0.21 | 0.24 | 0.19 | 0.2 | 0.21 | 0.2 | 0.27499 | 0.19 | 0.2 | 0.19761 | 0.2 | 0.2 | 0.21 | 0.18 | 0.3
| Mean |  | 0.21333 | 0.23 |  |  | 0.19833 |  |  | 0.2 | 0.14833 | 0.53333 | 0.26 | 0.20833 | 0.20167 | 0.20833 | 0.19333 | 0.20667 | 0.25863 | 0.19 | 0.21333 | 0.19562 | 0.21333 | 0.19 | 0.21833 | 0.18833 | 0.25
| Median |  | 0.215 | 0.23 |  |  | 0.2 |  |  | 0.2 | 0.15 | 0.5 | 0.225 | 0.21 | 0.2 | 0.205 | 0.195 | 0.205 | 0.23798 | 0.19 | 0.21 | 0.20308 | 0.215 | 0.195 | 0.22 | 0.185 | 0.25
| Std Dev. |  | 0.026583 | 0.02 |  |  | 0.0075277 |  |  | 3.0405e-17 | 0.018348 | 0.05164 | 0.095917 | 0.02137 | 0.01169 | 0.01472 | 0.012111 | 0.012111 | 0.044204 | 3.0405e-17 | 0.026583 | 0.027386 | 0.012111 | 0.016733 | 0.0075277 | 0.0098319 | 0.054772
| Rel.Std.Dev. |  | 12.46% | 8.696% |  |  | 3.795% |  |  | 1.52e-14% | 12.37% | 9.682% | 36.89% | 10.26% | 5.797% | 7.065% | 6.264% | 5.86% | 17.09% | 1.6e-14% | 12.46% | 14% | 5.677% | 8.807% | 3.448% | 5.22% | 21.91%
| PDM3 |  | 2.893% | 10.93% |  |  | -4.342% |  |  | -3.538% | -28.46% | 157.2% | 25.4% | 0.481% | -2.734% | 0.481% | -6.754% | -0.3228% | 24.74% | -8.361% | 2.893% | -5.651% | 2.893% | -8.361% | 5.304% | -9.165% | 20.58%
| Z-Score (Absolute) |  | 0.09 | 0.7 | 1.44 | 1.44 | 0.46 | 1.44 | 1.44 | NA | 2.3 | NA | 1.81 | 0.09 | 0.34 | 0.09 | 0.64 | 0.15 | 1.76 | 0.77 | 0.09 | 0.56 | 0.09 | 0.77 | 0.28 | 0.83 | NA
Results from laboratories 10, 12 & 30 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 554 (Certified Value 64 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 24 | 28 | 29 | 30 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 0
1
| 1 | 1 | 58.38 | 61.47 | 64.4 | 61.6 | 60 | 60 | 63.6 | 52.7 | 65.2 | 68.6 | 63.4 | 67.6 | 60.7 | 60.3 | 71.5366 | 64.52 | 65.4135 | 58.8 | 63.5 | 49.64 | 73 | [5 empty] | 1
| 1 | 2 | 58.01 | 62.12 | 63.6 | 59 | 58.5 | 59 | 64.1 | 50.9 | 66 | 69.5 | 66 | 68.7 | 63.6 | 61.4 | 71.9184 | 66.27 | 65.0248 | 60.15 | 63.19 | 50.55 | 71 | [5 empty] | 0
| 1 | 3 | 57.91 | 62.23 | 63.8 | 59.8 | 58.5 | 62 | 68.7 | 52 | 66.3 | 67.8 | 62.9 | 66.9 | 60 | 60.6 | 72.0308 | 70.3 | 66.8927 | 59.69 | 63.29 | 50.78 | 69 | [5 empty] | 16
| 1 | 4 | 56.98 | 62.62 | 62.4 | 59.3 | 62.5 | 63 | 64 | 50.2 | 65.9 | 68 | 63 | 68.1 | 63.6 | 61.6 | 70.5124 | 65.62 | 65.0528 | 60.05 | 62.95 | 51.36 | 73 | [5 empty] | 63.7732
| 1 | 5 | 56.88 | 61.87 | 63.5 | 60.2 | 60 | 60 | 66.5 | 50.9 | 63.2 | 69.5 | 64.1 | 67 | 61 | 61.7 | 70.907 | 65.43 | 65.3739 | 60.88 | 62.63 | 51.91 | 69 | [5 empty] | 90
| 1 | 6 | 58.4 | 61.85 | 65.2 | 60.3 | 57.5 | 59 | 63.7 | 50.9 | 62.9 | 65.3 | 66 | 66.9 | 59.4 | 60.6 | 71.1392 | 63.8 | 63.9821 | 60.54 | 64.52 | 49.48 | 73
| Mean |  | 57.76 | 62.0267 | 63.8167 | 60.0333 | 59.5 | 60.5 | 65.1 | 51.2667 | 64.9167 | 68.1167 | 64.2333 | 67.5333 | 61.3833 | 61.0333 | 71.3407 | 65.99 | 65.29 | 60.0183 | 63.3467 | 50.62 | 71.3333
| Median |  | 57.96 | 61.995 | 63.7 | 60 | 59.25 | 60 | 64.05 | 50.9 | 65.55 | 68.3 | 63.75 | 67.3 | 60.85 | 61 | 71.3379 | 65.525 | 65.2133 | 60.1 | 63.24 | 50.665 | 72
| Std Dev. |  | 0.67258 | 0.39185 | 0.93897 | 0.91797 | 1.7607 | 1.6432 | 2.0659 | 0.90921 | 1.4932 | 1.5562 | 1.432 | 0.74476 | 1.8049 | 0.60222 | 0.59382 | 2.2831 | 0.94106 | 0.72458 | 0.64791 | 0.94892 | 1.9664
| Rel.Std.Dev. |  | 1.164% | 0.6317% | 1.471% | 1.529% | 2.959% | 2.716% | 3.173% | 1.773% | 2.3% | 2.285% | 2.229% | 1.103% | 2.94% | 0.9867% | 0.8324% | 3.46% | 1.441% | 1.207% | 1.023% | 1.875% | 2.757%
| PDM3 |  | -9.429% | -2.739% | 0.0682% | -5.864% | -6.701% | -5.133% | 2.081% | -19.61% | 1.793% | 6.811% | 0.7216% | 5.896% | -3.747% | -4.296% | 11.87% | 3.476% | 2.378% | -5.888% | -0.6688% | -20.62% | 11.85%
| Z-Score (Absolute) |  | 1.32 | 0.31 | 0.11 | 0.78 | 0.91 | 0.67 | 0.42 | 2.86 | 0.37 | 1.13 | 0.21 | 0.99 | 0.47 | 0.55 | 1.89 | 0.63 | 0.46 | 0.79 | 0 | 3.01 | 1.89
| Analytical results for Co in OREAS 554 (Certified Value 0.815 wt.%) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 26 | 28 | 29 | 30 | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 3
3
| 1 | 1 | 0.7822 | 0.78936 | 0.802 | 0.82 | 0.80588 | 0.782 | 0.8808 | > 0.05 | 0.838 | 0.6928 | 0.836 | 0.831 | 0.813 | 0.831 | 0.786 | 0.88765 | 0.83334 | 0.839 | 0.83752 | 0.78722 | 0.798 | > 0.4 | 0.887 | 0.79995 | 0.7425 | 0.835 | 1
| 1 | 2 | 0.7787 | 0.80194 | 0.821 | 0.803 | 0.7921 | 0.7805 | 0.88 | > 0.05 | 0.831 | 0.6938 | 0.852 | 0.807 | 0.861 | 0.812 | 0.805 | 0.86422 | 0.9099 | 0.87 | 0.874 | 0.81823 | 0.796 | > 0.4 | 0.833 | 0.7962 | 0.7411 | 0.86 | 0
| 1 | 3 | 0.7861 | 0.81056 | 0.8 | 0.794 | 0.78846 | 0.7865 | 0.8652 | > 0.05 | 0.835 | 0.7053 | 0.844 | 0.797 | 0.837 | 0.802 | 0.807 | 0.91646 | 0.8829 | 0.847 | 0.83793 | 0.80224 | 0.799 | > 0.4 | 0.881 | 0.78046 | 0.7536 | 0.836 | 16
| 1 | 4 | 0.8054 | 0.81161 | 0.794 | 0.798 | 0.78747 | 0.789 | 0.8574 | > 0.05 | 0.822 | 0.681 | 0.812 | 0.78 | 0.809 | 0.817 | 0.75 | 0.87526 | 0.84842 | 0.844 | 0.86665 | 0.79274 | 0.793 | > 0.4 | 0.834 | 0.7764 | 0.7533 | 0.842 | 0.81545
| 1 | 5 | 0.7944 | 0.79481 | 0.786 | 0.796 | 0.80295 | 0.7775 | 0.8459 | > 0.05 | 0.851 | 0.6955 | 0.798 | 0.808 | 0.818 | 0.798 | 0.787 | 0.8964 | 0.81548 | 0.855 | 0.89299 | 0.806 | 0.798 | > 0.4 | 0.881 | 0.79585 | 0.7662 | 0.834 | 91
| 1 | 6 | 0.8053 | 0.79949 | 0.801 | 0.795 | 0.80085 | 0.788 | 0.8527 | > 0.05 | 0.819 | 0.6904 | 0.81 | 0.839 | 0.822 | 0.837 | 0.768 | 0.87192 | 0.82754 | 0.844 | 0.8816 | 0.80939 | 0.8 | > 0.4 | 0.868 | 0.77172 | 0.7613 | 0.877
| Mean |  | 0.79202 | 0.80129 | 0.80067 | 0.801 | 0.79629 | 0.78392 | 0.86367 |  | 0.83267 | 0.69313 | 0.82533 | 0.81033 | 0.82667 | 0.81617 | 0.78383 | 0.88532 | 0.85293 | 0.84983 | 0.86511 | 0.80264 | 0.79733 |  | 0.864 | 0.78676 | 0.753 | 0.84733
| Median |  | 0.79025 | 0.80071 | 0.8005 | 0.797 | 0.79648 | 0.78425 | 0.8613 |  | 0.833 | 0.6933 | 0.824 | 0.8075 | 0.82 | 0.8145 | 0.7865 | 0.88145 | 0.84088 | 0.8455 | 0.87033 | 0.80412 | 0.798 |  | 0.8745 | 0.78816 | 0.75345 | 0.839
| Std Dev. |  | 0.011575 | 0.0087225 | 0.011622 | 0.0098387 | 0.0079218 | 0.004598 | 0.014407 |  | 0.011605 | 0.0078597 | 0.021602 | 0.021741 | 0.019398 | 0.01551 | 0.021885 | 0.019099 | 0.036318 | 0.011197 | 0.022937 | 0.011276 | 0.0025033 |  | 0.024429 | 0.011991 | 0.0099543 | 0.017478
| Rel.Std.Dev. |  | 1.462% | 1.089% | 1.452% | 1.228% | 0.9948% | 0.5865% | 1.668% |  | 1.394% | 1.134% | 2.617% | 2.683% | 2.346% | 1.9% | 2.792% | 2.157% | 4.258% | 1.318% | 2.651% | 1.405% | 0.314% |  | 2.827% | 1.524% | 1.322% | 2.063%
| PDM3 |  | -2.874% | -1.736% | -1.813% | -1.773% | -2.351% | -3.867% | 5.912% |  | 2.111% | -15% | 1.211% | -0.628% | 1.375% | 0.08736% | -3.878% | 8.567% | 4.596% | 4.216% | 6.09% | -1.572% | -2.222% |  | 5.953% | -3.518% | -7.659% | 3.909%
| Z-Score (Absolute) |  | 0.46 | 0.16 | 0.18 | 0.17 | 0.32 | 0.71 | 1.8 | NA | 0.83 | 3.58 | 0.59 | 0.12 | 0.64 | 0.31 | 0.71 | 2.49 | 1.47 | 1.37 | 1.85 | 0.12 | 0.29 | NA | 1.81 | 0.62 | 1.69 | 1.29
| Analytical results for Cr in OREAS 554 (Certified Value 49.5 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
1
| 1 | 1 | 54 | 53 | 30 | 50 | 41 | 60 | 45 | 49 | 27 | 48 | 53 | 48 | 49 | 47 | 51.8 | 27.1732 | 23 | 61.4348 | 50 | 75 | 49 | 48 | 48 |  |  |  | 1
| 1 | 2 | 47 | 58 | 30 | 40 | 40 | 60 | 43 | 47 | 28 | 51 | 56 | 49 | 49 | 48 | 53.2 | 28.8994 | 26 | 64.1001 | 50 | 71 | 50 | 49 | 52 |  |  |  | 0
| 1 | 3 | 54 | 57 | 30 | 40 | 41 | 60 | 42 | 56 | 28 | 50 | 52 | 50 | 48 | 47 | 55.9 | 25.8039 | 77 | 64.3224 | 50 | 71 | 47 | 51 | 54 |  |  |  | 16
| 1 | 4 | 57 | 65 | 40 | 40 | 41 | 45 | 40 | 49 | 27 | 50 | 53 | 50 | 47 | 46 | 50.6 | 23.9994 | 36 | 63.3953 | 49 | 73 | 47 | 50 | 47 |  |  |  | 49.4667
| 1 | 5 | 52 | 58 | 60 | 40 | 40 | 45 | 42 | 44 | 27 | 47 | 47 | 53 | 46 | 49 | 52.6 | 27.811 | 34 | 61.7433 | 50 | 70 | 52 | 56 | 50 |  |  |  | 92
| 1 | 6 | 52 | 59 | 60 | 40 | 42 | 60 | 42 | 47 | 28 | 48 | 49 | 52 | 49 | 47 | 50.5 | 28.5886 | 65 | 59.6482 | 49 | 85 | 43 | 53 | 52
| Mean |  | 52.6667 | 58.3333 | 41.6667 | 41.6667 | 40.8333 | 55 | 42.3333 | 48.6667 | 27.5 | 49 | 51.6667 | 50.3333 | 48 | 47.3333 | 52.4333 | 27.0459 | 43.5 | 62.4407 | 49.6667 | 74.1667 | 48 | 51.1667 | 50.5
| Median |  | 53 | 58 | 35 | 40 | 41 | 60 | 42 | 48 | 27.5 | 49 | 52.5 | 50 | 48.5 | 47 | 52.2 | 27.4921 | 35 | 62.5693 | 50 | 72 | 48 | 50.5 | 51
| Std Dev. |  | 3.3267 | 3.8816 | 14.7196 | 4.0825 | 0.75277 | 7.746 | 1.633 | 4.0332 | 0.54772 | 1.5492 | 3.2042 | 1.8619 | 1.2649 | 1.0328 | 2.0067 | 1.8569 | 22.1698 | 1.8167 | 0.5164 | 5.6006 | 3.0984 | 2.9269 | 2.6646
| Rel.Std.Dev. |  | 6.316% | 6.654% | 35.33% | 9.798% | 1.844% | 14.08% | 3.857% | 8.287% | 1.992% | 3.162% | 6.202% | 3.699% | 2.635% | 2.182% | 3.827% | 6.866% | 50.97% | 2.909% | 1.04% | 7.551% | 6.455% | 5.72% | 5.276%
| PDM3 |  | 6.469% | 17.92% | -15.77% | -15.77% | -17.45% | 11.19% | -14.42% | -1.617% | -44.41% | -0.9434% | 4.447% | 1.752% | -2.965% | -4.313% | 5.997% | -45.32% | -12.06% | 26.23% | 0.4043% | 49.93% | -2.965% | 3.437% | 2.089%
| Z-Score (Absolute) |  | 0.73 | 2.11 | NA | NA | 2.15 | 1.3 | 1.79 | 0.24 | 5.4 | 0.16 | 0.49 | 0.16 | 0.41 | 0.57 | 0.67 | 5.51 | 1.5 | 3.11 | 0 | 5.97 | 0.41 | 0.37 | 0.2
Results from laboratories 3 & 4 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 554 (Certified Value 2.42 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 3 | 2.52 | 2.4 | 2.3 | 2.4 | 2.5 | 2.28 | 2.52 | 2.61 | 2.61 | 2.38 | 2.31 | 2.46 | 2.5 | 2.8741 | < 5 | 2.3304 | 2.4 | 2.3 | 2.51 | 2.39 | 2.3 |  |  |  |  | 1
| 1 | 2 | 2 | 2.58 | 2.3 | 2.3 | 2.4 | 2.5 | 2.3 | 2.58 | 2.68 | 2.68 | 2.43 | 2.31 | 2.52 | 2.5 | 2.8995 | < 5 | 2.3519 | 2.3 | 2.2 | 2.54 | 2.37 | 2.5 |  |  |  |  | 0
| 1 | 3 | 2 | 2.54 | 2.3 | 2.3 | 2.4 | 2.5 | 2.54 | 2.58 | 2.61 | 2.57 | 2.4 | 2.27 | 2.4 | 2.5 | 2.9202 | < 5 | 2.3421 | 2.3 | 2.2 | 2.53 | 2.42 | 2.4 |  |  |  |  | 16
| 1 | 4 | 3 | 2.51 | 2.3 | 2.3 | 2.4 | 2.5 | 2.33 | 2.61 | 2.54 | 2.58 | 2.38 | 2.32 | 2.49 | 2.5 | 2.8118 | < 5 | 2.3818 | 2.3 | 2.3 | 2.54 | 2.47 | 2.4 |  |  |  |  | 2.4249
| 1 | 5 | 2 | 2.58 | 2.4 | 2.3 | 2.4 | 2.5 | 2.38 | 2.65 | 2.41 | 2.6 | 2.38 | 2.28 | 2.42 | 2.5 | 2.7983 | < 5 | 2.3513 | 2.4 | 2.1 | 2.53 | 2.46 | 2.3 |  |  |  |  | 93
| 1 | 6 | 3 | 2.44 | 2.4 | 2.3 | 2.4 | 2.5 | 2.37 | 2.5 | 2.42 | 2.61 | 2.44 | 2.32 | 2.4 | 2.5 | 2.8303 | < 5 | 2.3708 | 2.3 | 2.3 | 2.5 | 2.41 | 2.4
| Mean |  | 2.5 | 2.5283 | 2.35 | 2.3 | 2.4 | 2.5 | 2.3667 | 2.5733 | 2.545 | 2.6083 | 2.4017 | 2.3017 | 2.4483 | 2.5 | 2.8557 |  | 2.3547 | 2.3333 | 2.2333 | 2.525 | 2.42 | 2.3833
| Median |  | 2.5 | 2.53 | 2.35 | 2.3 | 2.4 | 2.5 | 2.35 | 2.58 | 2.575 | 2.605 | 2.39 | 2.31 | 2.44 | 2.5 | 2.8522 |  | 2.3516 | 2.3 | 2.25 | 2.53 | 2.415 | 2.4
| Std Dev. |  | 0.54772 | 0.052313 | 0.054772 | 4.8648e-16 | 0 | 0 | 0.09331 | 0.055737 | 0.11005 | 0.038687 | 0.027142 | 0.02137 | 0.049967 | 0 | 0.049575 |  | 0.01879 | 0.05164 | 0.08165 | 0.016432 | 0.038987 | 0.075277
| Rel.Std.Dev. |  | 21.91% | 2.069% | 2.331% | 2.115e-14% | 0% | 0% | 3.943% | 2.166% | 4.324% | 1.483% | 1.13% | 0.9284% | 2.041% | 0% | 1.736% |  | 0.798% | 2.213% | 3.656% | 0.6508% | 1.611% | 3.158%
| PDM3 |  | 3.097% | 4.266% | -3.089% | -5.151% | -1.027% | 3.097% | -2.401% | 6.121% | 4.953% | 7.565% | -0.958% | -5.082% | 0.9665% | 3.097% | 17.77% |  | -2.894% | -3.776% | -7.9% | 4.128% | -0.202% | -1.714%
| Z-Score (Absolute) |  | NA | 0.91 | 0.59 | 1.01 | 0.17 | 0.67 | 0.45 | 1.29 | 1.05 | 1.59 | 0.15 | 1 | 0.24 | 0.67 | 3.67 | 0.67 | 0.55 | 0.73 | 1.57 | 0.89 | 0 | 0.31
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 554 (Certified Value 1.59 wt.%) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 26 | 28 | 29 | 30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*AAS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*AAS | 4A*MS |  | 2
3
| 1 | 1 | 1.52 | 1.5596 | 1.58 | 1.6 | 1.6202 | 1.55 | 1.5321 | > 1 | 1.72 | 1.61 | 1.595 | 1.585 | 1.58 | 1.595 | 1.5742 | 1.533 | 1.63 | 1.5533 | 1.5105 | 1.61 | > 1 | 1.739 | > 1 | 1.586 | 1.58 |  | 1
| 1 | 2 | 1.5 | 1.6031 | 1.61 | 1.56 | 1.5728 | 1.563 | 1.5496 | > 1 | 1.73 | 1.56 | 1.645 | 1.605 | 1.595 | 1.575 | 1.566 | 1.503 | 1.66 | 1.5656 | 1.5701 | 1.6 | > 1 | 1.715 | > 1 | 1.564 | 1.64 |  | 0
| 1 | 3 | 1.53 | 1.577 | 1.58 | 1.55 | 1.5901 | 1.551 | 1.5511 | > 1 | 1.71 | 1.605 | 1.59 | 1.595 | 1.585 | 1.61 | 1.5905 | 1.539 | 1.65 | 1.5904 | 1.5784 | 1.59 | > 1 | 1.663 | > 1 | 1.576 | 1.61 |  | 16
| 1 | 4 | 1.46 | 1.5881 | 1.55 | 1.61 | 1.5752 | 1.54 | 1.5299 | > 1 | 1.71 | 1.59 | 1.59 | 1.595 | 1.59 | 1.61 | 1.5756 | 1.518 | 1.61 | 1.5879 | 1.57 | 1.58 | > 1 | 1.731 | > 1 | 1.558 | 1.62 |  | 1.5881
| 1 | 5 | 1.53 | 1.5964 | 1.55 | 1.57 | 1.6088 | 1.563 | 1.5271 | > 1 | 1.72 | 1.575 | 1.57 | 1.57 | 1.58 | 1.595 | 1.5946 | 1.51 | 1.6 | 1.6298 | 1.5767 | 1.63 | > 1 | 1.64 | > 1 | 1.574 | 1.62 |  | 94
| 1 | 6 | 1.46 | 1.563 | 1.59 | 1.56 | 1.6207 | 1.559 | 1.5277 | > 1 | 1.72 | 1.6 | 1.555 | 1.595 | 1.59 | 1.555 | 1.5914 | 1.524 | 1.63 | 1.6088 | 1.5807 | 1.62 | > 1 | 1.652 | > 1 | 1.571 | 1.6
| Mean |  | 1.5 | 1.5812 | 1.5767 | 1.575 | 1.598 | 1.5543 | 1.5363 |  | 1.7183 | 1.59 | 1.5908 | 1.5908 | 1.5867 | 1.59 | 1.582 | 1.5212 | 1.63 | 1.5893 | 1.5644 | 1.605 |  | 1.69 |  | 1.5715 | 1.6117
| Median |  | 1.51 | 1.5825 | 1.58 | 1.565 | 1.5995 | 1.555 | 1.531 |  | 1.72 | 1.595 | 1.59 | 1.595 | 1.5875 | 1.595 | 1.5831 | 1.521 | 1.63 | 1.5892 | 1.5734 | 1.605 |  | 1.689 |  | 1.5725 | 1.615
| Std Dev. |  | 0.032863 | 0.017749 | 0.023381 | 0.02429 | 0.021643 | 0.0090259 | 0.011086 |  | 0.0075277 | 0.019235 | 0.030564 | 0.012007 | 0.0060553 | 0.021448 | 0.011642 | 0.013644 | 0.022804 | 0.027862 | 0.026758 | 0.018708 |  | 0.043313 |  | 0.0097519 | 0.020412
| Rel.Std.Dev. |  | 2.191% | 1.123% | 1.483% | 1.542% | 1.354% | 0.5807% | 0.7216% |  | 0.4381% | 1.21% | 1.921% | 0.7548% | 0.3816% | 1.349% | 0.7359% | 0.897% | 1.399% | 1.753% | 1.71% | 1.166% |  | 2.563% |  | 0.6205% | 1.267%
| PDM3 |  | -5.549% | -0.4372% | -0.7216% | -0.8265% | 0.6196% | -2.128% | -3.266% |  | 8.199% | 0.118% | 0.1704% | 0.1704% | -0.09193% | 0.118% | -0.3836% | -4.216% | 2.637% | 0.07388% | -1.495% | 1.062% |  | 6.415% |  | -1.047% | 1.482%
| Z-Score (Absolute) |  | 4.03 | 0.31 | 0.52 | 0.59 | 0.46 | 1.54 | 2.37 | NA | 5.97 | 0.09 | 0.13 | 0.13 | 0.06 | 0.09 | 0.27 | 3.06 | 1.92 | 0.06 | 1.08 | 0.78 | NA | 4.67 | NA | 0.75 | 1.08
| Analytical results for Dy in OREAS 554 (Certified Value 2.71 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 22; 12: 24; 13: 29; 14: 30 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*MS |  | 2
2
| 1 | 1 | 01: 2.57; 02: 2.59; 03: 3; 04: 2.7; 05: 2.5; 06: 2.9; 09: 2.85; 10: 3.1504; 12: 2.4; 13: 2.47; 14: 2.58 |  | 1
| 1 | 2 | 01: 2.57; 02: 2.63; 03: 2.95; 04: 2.7; 05: 2.3; 06: 3; 09: 2.86; 10: 3.3497; 12: 2.3; 13: 2.37; 14: 3.14 |  | 0
| 1 | 3 | 01: 2.5; 02: 2.7; 03: 2.8; 04: 2.8; 05: 2.4; 06: 3.1; 09: 2.71; 10: 3.3008; 12: 2.4; 13: 2.37; 14: 2.52 |  | 16
| 1 | 4 | 01: 2.54; 02: 2.71; 03: 2.9; 04: 2.75; 05: 2.7; 06: 3.1; 09: 2.84; 10: 3.354; 12: 2.4; 13: 2.39; 14: 2.5 |  | 2.7091
| 1 | 5 | 01: 2.54; 02: 2.6; 03: 2.9; 04: 2.75; 05: 2.5; 06: 3.1; 09: 3; 10: 3.3364; 12: 2.3; 13: 2.45; 14: 2.6 |  | 95
| 1 | 6 | 01: 2.59; 02: 2.6; 03: 2.9; 04: 2.75; 05: 2.3; 06: 3; 09: 2.67; 10: 3.2846; 12: 2.5; 13: 2.38; 14: 2.51
| Mean |  | 01: 2.5517; 02: 2.6383; 03: 2.9083; 04: 2.7417; 05: 2.45; 06: 3.0333; 09: 2.8217; 10: 3.296; 12: 2.3833; 13: 2.405; 14: 2.6417
| Median |  | 01: 2.555; 02: 2.615; 03: 2.9; 04: 2.75; 05: 2.45; 06: 3.05; 09: 2.845; 10: 3.3186; 12: 2.4; 13: 2.385; 14: 2.55
| Std Dev. |  | 01: 0.031885; 02: 0.053448; 03: 0.066458; 04: 0.037639; 05: 0.15166; 06: 0.08165; 09: 0.11822; 10: 0.076449; 12: 0.075277; 13: 0.043704; 14: 0.24742
| Rel.Std.Dev. |  | 01: 1.25%; 02: 2.026%; 03: 2.285%; 04: 1.373%; 05: 6.19%; 06: 2.692%; 09: 4.19%; 10: 2.319%; 12: 3.158%; 13: 1.817%; 14: 9.366%
| PDM3 |  | 01: -5.812%; 02: -2.613%; 03: 7.353%; 04: 1.201%; 05: -9.565%; 06: 11.97%; 09: 4.154%; 10: 21.66%; 12: -12.03%; 13: -11.23%; 14: -2.49%
| Z-Score (Absolute) |  | 01: 0.32; 02: 0.01; 03: 0.94; 04: 0.35; 05: 0.67; 06: 1.38; 09: 0.63; 10: 2.3; 12: 0.91; 13: 0.83; 14: 0
| Analytical results for Er in OREAS 554 (Certified Value 1.43 ppm) | [27 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 22; 12: 24; 13: 29; 14: 30 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*MS |  | 2
2
| 1 | 1 | 01: 1.46; 02: 1.28; 03: 1.6; 04: 1.45; 05: 1.15; 06: 1.6; 09: 1.38; 10: 1.9255; 12: 1.2; 13: 1.18; 14: 1.22 |  | 1
| 1 | 2 | 01: 1.46; 02: 1.32; 03: 1.55; 04: 1.55; 05: 1.25; 06: 1.7; 09: 1.48; 10: 1.9056; 12: 1.1; 13: 1.21; 14: 1.22 |  | 0
| 1 | 3 | 01: 1.43; 02: 1.33; 03: 1.55; 04: 1.5; 05: 1.3; 06: 1.8; 09: 1.44; 10: 1.9051; 12: 1.1; 13: 1.19; 14: 1.21 |  | 16
| 1 | 4 | 01: 1.42; 02: 1.28; 03: 1.6; 04: 1.55; 05: 1.35; 06: 1.7; 09: 1.45; 10: 1.9717; 12: 1.1; 13: 1.28; 14: 1.21 |  | 1.4316
| 1 | 5 | 01: 1.39; 02: 1.32; 03: 1.6; 04: 1.45; 05: 1.3; 06: 1.7; 09: 1.31; 10: 1.9138; 12: 1.4; 13: 1.29; 14: 1.19 |  | 96
| 1 | 6 | 01: 1.45; 02: 1.34; 03: 1.55; 04: 1.45; 05: 1.25; 06: 1.6; 09: 1.5; 10: 1.9732; 12: 1.1; 13: 1.26; 14: 1.21
| Mean |  | 01: 1.435; 02: 1.3117; 03: 1.575; 04: 1.4917; 05: 1.2667; 06: 1.6833; 09: 1.4267; 10: 1.9325; 12: 1.1667; 13: 1.235; 14: 1.21
| Median |  | 01: 1.44; 02: 1.32; 03: 1.575; 04: 1.475; 05: 1.275; 06: 1.7; 09: 1.445; 10: 1.9196; 12: 1.1; 13: 1.235; 14: 1.21
| Std Dev. |  | 01: 0.027386; 02: 0.025626; 03: 0.027386; 04: 0.04916; 05: 0.068313; 06: 0.075277; 09: 0.070333; 10: 0.031819; 12: 0.12111; 13: 0.047645; 14: 0.010954
| Rel.Std.Dev. |  | 01: 1.908%; 02: 1.954%; 03: 1.739%; 04: 3.296%; 05: 5.393%; 06: 4.472%; 09: 4.93%; 10: 1.647%; 12: 10.38%; 13: 3.858%; 14: 0.9053%
| PDM3 |  | 01: 0.2396%; 02: -8.376%; 03: 10.02%; 04: 4.198%; 05: -11.52%; 06: 17.59%; 09: -0.3425%; 10: 34.99%; 12: -18.5%; 13: -13.73%; 14: -15.48%
| Z-Score (Absolute) |  | 01: 0.04; 02: 0.52; 03: 0.67; 04: 0.3; 05: 0.73; 06: NA; 09: 0; 10: 2.3; 12: NA; 13: 0.87; 14: 0.98
Results from laboratories 10 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 554 (Certified Value 0.82 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 12: Lab; 13: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 22; 10: 24; 12: 29; 13: 30 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 12: 4A*OES/MS; 13: 4A*MS |  | 2
3
| 1 | 1 | 01: 1.15; 02: 0.82; 03: 0.8; 04: 0.8; 05: 0.6; 06: 0.84; 09: 0.7828; 10: 0.8; 12: 1.16; 13: 0.85 |  | 1
| 1 | 2 | 01: 1.24; 02: 0.85; 03: 0.8; 04: 0.8; 05: 0.6; 06: 0.83; 09: 0.82679; 10: 0.8; 12: 1.17; 13: 0.9 |  | 0
| 1 | 3 | 01: 1.21; 02: 0.82; 03: 0.75; 04: 0.8; 05: 0.5; 06: 0.88; 09: 0.81485; 10: 0.8; 12: 1.2; 13: 0.86 |  | 16
| 1 | 4 | 01: 1.23; 02: 0.84; 03: 0.8; 04: 0.8; 05: 0.6; 06: 0.86; 09: 0.81935; 10: 0.8; 12: 1.18; 13: 0.84 |  | 0.8243
| 1 | 5 | 01: 1.2; 02: 0.82; 03: 0.8; 04: 0.75; 05: 0.6; 06: 0.87; 09: 0.80484; 10: 0.8; 12: 1.22; 13: 0.85 |  | 97
| 1 | 6 | 01: 1.2; 02: 0.79; 03: 0.75; 04: 0.8; 05: 0.6; 06: 0.84; 09: 0.80839; 10: 0.8; 12: 1.21; 13: 0.86
| Mean |  | 01: 1.205; 02: 0.82333; 03: 0.78333; 04: 0.79167; 05: 0.58333; 06: 0.85333; 09: 0.8095; 10: 0.8; 12: 1.19; 13: 0.86
| Median |  | 01: 1.205; 02: 0.82; 03: 0.8; 04: 0.8; 05: 0.6; 06: 0.85; 09: 0.81162; 10: 0.8; 12: 1.19; 13: 0.855
| Std Dev. |  | 01: 0.031464; 02: 0.020656; 03: 0.02582; 04: 0.020412; 05: 0.040825; 06: 0.019664; 09: 0.015234; 10: 1.2162e-16; 12: 0.023664; 13: 0.020976
| Rel.Std.Dev. |  | 01: 2.611%; 02: 2.509%; 03: 3.296%; 04: 2.578%; 05: 6.999%; 06: 2.304%; 09: 1.882%; 10: 1.52e-14%; 12: 1.989%; 13: 2.439%
| PDM3 |  | 01: 46.18%; 02: -0.1174%; 03: -4.97%; 04: -3.959%; 05: -29.23%; 06: 3.522%; 09: -1.795%; 10: -2.948%; 12: 44.36%; 13: 4.331%
| Z-Score (Absolute) |  | 01: 5.41; 02: 0.46; 03: 1.08; 04: NA; 05: NA; 06: 0; 09: 0.67; 10: NA; 12: 5.18; 13: 0.1
Results from laboratories 4, 6 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 554 (Certified Value 1.8 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 1.88 | 1.7 | 1.83 | 1.85 | 1.93 | 1.835 | 1.81 | 1.85 | 2.04 | 1.81 | 1.76 | 1.82 | 1.72 | 1.79 | 1.74 | 1.7578 | 1.8469 | 1.43 | 1.77 | 1.7 | 1.8156 | 1.81 | 1.76 | 1.909 |  |  | 1
| 1 | 2 | 1.87 | 1.68 | 1.82 | 1.84 | 1.87 | 1.82 | 1.81 | 1.83 | 2.04 | 1.84 | 1.81 | 1.79 | 1.8 | 1.78 | 1.71 | 1.7692 | 1.9275 | 1.43 | 1.8 | 1.73 | 1.8564 | 1.81 | 1.77 | 1.937 |  |  | 0
| 1 | 3 | 1.88 | 1.69 | 1.85 | 1.85 | 1.86 | 1.825 | 1.81 | 1.99 | 2.08 | 1.84 | 1.79 | 1.76 | 1.77 | 1.74 | 1.68 | 1.8098 | 1.744 | 1.45 | 1.82 | 1.66 | 1.8156 | 1.8 | 1.78 | 1.951 |  |  | 16
| 1 | 4 | 1.89 | 1.72 | 1.78 | 1.85 | 1.91 | 1.82 | 1.79 | 1.87 | 2.08 | 1.81 | 1.76 | 1.78 | 1.75 | 1.78 | 1.65 | 1.7445 | 1.8757 | 1.39 | 1.79 | 1.63 | 1.8972 | 1.84 | 1.73 | 1.937 |  |  | 1.8036
| 1 | 5 | 1.87 | 1.71 | 1.82 | 1.83 | 1.92 | 1.81 | 1.77 | 1.88 | 2.04 | 1.84 | 1.72 | 1.78 | 1.77 | 1.75 | 1.72 | 1.7776 | 1.7802 | 1.44 | 1.83 | 1.67 | 1.8462 | 1.84 | 1.78 | 1.937 |  |  | 98
| 1 | 6 | 1.87 | 1.69 | 1.83 | 1.87 | 1.93 | 1.83 | 1.78 | 1.84 | 2.09 | 1.83 | 1.73 | 1.79 | 1.79 | 1.82 | 1.67 | 1.7562 | 1.7952 | 1.45 | 1.78 | 1.7 | 1.9278 | 1.77 | 1.74 | 1.937
| Mean |  | 1.8767 | 1.6983 | 1.8217 | 1.8483 | 1.9033 | 1.8233 | 1.795 | 1.8767 | 2.0617 | 1.8283 | 1.7617 | 1.7867 | 1.7667 | 1.7767 | 1.695 | 1.7692 | 1.8282 | 1.4317 | 1.7983 | 1.6817 | 1.8598 | 1.8117 | 1.76 | 1.9347
| Median |  | 1.875 | 1.695 | 1.825 | 1.85 | 1.915 | 1.8225 | 1.8 | 1.86 | 2.06 | 1.835 | 1.76 | 1.785 | 1.77 | 1.78 | 1.695 | 1.7635 | 1.821 | 1.435 | 1.795 | 1.685 | 1.8513 | 1.81 | 1.765 | 1.937
| Std Dev. |  | 0.008165 | 0.01472 | 0.023166 | 0.013292 | 0.030768 | 0.008756 | 0.017607 | 0.058538 | 0.024014 | 0.01472 | 0.034303 | 0.019664 | 0.028752 | 0.028752 | 0.033912 | 0.022926 | 0.067731 | 0.022286 | 0.023166 | 0.035449 | 0.045003 | 0.026394 | 0.020976 | 0.013765
| Rel.Std.Dev. |  | 0.4351% | 0.8667% | 1.272% | 0.7191% | 1.617% | 0.4802% | 0.9809% | 3.119% | 1.165% | 0.8051% | 1.947% | 1.101% | 1.627% | 1.618% | 2.001% | 1.296% | 3.705% | 1.557% | 1.288% | 2.108% | 2.42% | 1.457% | 1.192% | 0.7115%
| PDM3 |  | 4.05% | -5.838% | 1% | 2.479% | 5.528% | 1.093% | -0.478% | 4.05% | 14.31% | 1.37% | -2.326% | -0.9401% | -2.049% | -1.495% | -6.022% | -1.909% | 1.364% | -20.62% | -0.2932% | -6.762% | 3.115% | 0.446% | -2.419% | 7.266%
| Z-Score (Absolute) |  | 1.12 | 1.66 | 0.26 | 0.67 | 1.53 | 0.29 | 0.16 | 1.12 | 3.99 | 0.36 | 0.67 | 0.29 | 0.6 | 0.44 | 1.71 | 0.56 | 0.36 | 5.81 | 0.1 | 1.92 | 0.85 | 0.1 | 0.7 | 2.03
| Analytical results for Ga in OREAS 554 (Certified Value 12.5 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 12.5 | 11.94 | 13 | 13.2 | 12.6 | 15 | 11.4 | 12.5 | 13 | 12.75 | 13.05 | 12.85 | 12.55 | 12.75 | 13.19 | 12.068 | 12.2 | 12.5 | 12.2365 | 13.9 | 13.12 | 12.57 | 11.9 | 10 |  |  | 1
| 1 | 2 | 12.2 | 11.7 | 13 | 12.8 | 12.2 | 15 | 11 | 12.3 | 14 | 12.8 | 13 | 13.3 | 12.55 | 13 | 13.27 | 12.5066 | 11.9 | 12.9 | 12.3338 | 13.3 | 11.45 | 12.45 | 11.7 | 10.6 |  |  | 0
| 1 | 3 | 12.2 | 12.1 | 13 | 13 | 12.2 | 15 | 12 | 12.2 | 13 | 12.6 | 12.75 | 12.8 | 12.4 | 12.6 | 13.32 | 12.1409 | 12.2 | 13.5 | 12.1268 | 13.3 | 11.87 | 12.51 | 11.8 | 10.4 |  |  | 16
| 1 | 4 | 12.4 | 11.67 | 12.8 | 13 | 12.6 | 15 | 11.2 | 13.1 | 12 | 12.35 | 12.9 | 12.75 | 12.6 | 12.7 | 13.02 | 12.0007 | 12.1 | 13.2 | 12.3169 | 13.3 | 11.29 | 12.44 | 12.1 | 10.2 |  |  | 12.5157
| 1 | 5 | 12 | 11.3 | 13.2 | 13 | 12.4 | 15 | 11.5 | 13.4 | 13 | 11.65 | 13.15 | 12.7 | 12.45 | 12.65 | 13.44 | 11.8393 | 12.1 | 13.7 | 12.2734 | 13.9 | 11.9 | 12.53 | 12.2 | 10.4 |  |  | 99
| 1 | 6 | 12.2 | 11.84 | 13.2 | 13 | 12.5 | 10 | 11 | 12.4 | 13 | 11.75 | 13.75 | 13.05 | 12.5 | 12.35 | 13.37 | 11.8543 | 12.2 | 13.1 | 12.4354 | 14.1 | 12.36 | 12.42 | 12.2 | 10.4
| Mean |  | 12.25 | 11.7583 | 13.0333 | 13 | 12.4167 | 14.1667 | 11.35 | 12.65 | 13 | 12.3167 | 13.1 | 12.9083 | 12.5083 | 12.675 | 13.2683 | 12.0683 | 12.1167 | 13.15 | 12.2871 | 13.6333 | 11.9983 | 12.4867 | 11.9833 | 10.3333
| Median |  | 12.2 | 11.77 | 13 | 13 | 12.45 | 15 | 11.3 | 12.45 | 13 | 12.475 | 13.025 | 12.825 | 12.525 | 12.675 | 13.295 | 12.0343 | 12.15 | 13.15 | 12.2951 | 13.6 | 11.885 | 12.48 | 12 | 10.4
| Std Dev. |  | 0.17607 | 0.27484 | 0.15055 | 0.12649 | 0.18348 | 2.0412 | 0.37815 | 0.48477 | 0.63246 | 0.50365 | 0.34641 | 0.22675 | 0.073598 | 0.21154 | 0.14851 | 0.24503 | 0.1169 | 0.42778 | 0.10346 | 0.37238 | 0.66596 | 0.058878 | 0.2137 | 0.20656
| Rel.Std.Dev. |  | 1.437% | 2.337% | 1.155% | 0.973% | 1.478% | 14.41% | 3.332% | 3.832% | 4.865% | 4.089% | 2.644% | 1.757% | 0.5884% | 1.669% | 1.119% | 2.03% | 0.9648% | 3.253% | 0.842% | 2.731% | 5.55% | 0.4715% | 1.783% | 1.999%
| PDM3 |  | -2.123% | -6.051% | 4.136% | 3.87% | -0.7909% | 13.19% | -9.314% | 1.073% | 3.87% | -1.59% | 4.669% | 3.138% | -0.0585% | 1.273% | 6.014% | -3.575% | -3.188% | 5.068% | -1.826% | 8.93% | -4.133% | -0.2316% | -4.253% | -17.44%
| Z-Score (Absolute) |  | 0.3 | 1.03 | 0.86 | 0.81 | 0.05 | NA | 1.63 | 0.29 | NA | 0.2 | 0.96 | 0.68 | 0.08 | 0.33 | 1.21 | 0.57 | 0.5 | 1.03 | 0.24 | 1.75 | 0.67 | 0.05 | 0.69 | 3.14
Results from laboratories 6 & 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 554 (Certified Value 3.78 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 10: 11; 11: 22; 12: 24; 13: 29; 14: 30 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*MS |  |  | 2
3
| 1 | 1 | 01: 4.2; 02: 3.67; 03: 3.8; 04: 3.8; 05: 3.8; 06: 3.9; 10: 4.72; 11: 3.6831; 12: 3.3; 13: 3.7; 14: 3.3 |  |  | 1
| 1 | 2 | 01: 4.1; 02: 3.69; 03: 4; 04: 3.8; 05: 3.8; 06: 3.9; 10: 4.44; 11: 3.875; 12: 3.4; 13: 3.71; 14: 3.4 |  |  | 0
| 1 | 3 | 01: 4.1; 02: 3.62; 03: 3.8; 04: 3.6; 05: 3.6; 06: 4.2; 10: 4.91; 11: 3.7728; 12: 3.1; 13: 3.66; 14: 3.3 |  |  | 16
| 1 | 4 | 01: 4.1; 02: 3.67; 03: 4; 04: 3.8; 05: 3.8; 06: 3.9; 10: 5.31; 11: 3.7722; 12: 3.4; 13: 3.86; 14: 3.3 |  |  | 3.7817
| 1 | 5 | 01: 4.2; 02: 3.63; 03: 3.8; 04: 3.8; 05: 3.8; 06: 3.9; 10: 5.31; 11: 4.0057; 12: 3.4; 13: 3.74; 14: 3.2 |  |  | 100
| 1 | 6 | 01: 4.1; 02: 3.6; 03: 3.8; 04: 3.8; 05: 3.5; 06: 3.9; 10: 5.19; 11: 3.8903; 12: 3.5; 13: 3.7; 14: 3.3
| Mean |  | 01: 4.1333; 02: 3.6467; 03: 3.8667; 04: 3.7667; 05: 3.7167; 06: 3.95; 10: 4.98; 11: 3.8332; 12: 3.35; 13: 3.7283; 14: 3.3
| Median |  | 01: 4.1; 02: 3.65; 03: 3.8; 04: 3.8; 05: 3.8; 06: 3.9; 10: 5.05; 11: 3.8239; 12: 3.4; 13: 3.705; 14: 3.3
| Std Dev. |  | 01: 0.05164; 02: 0.035024; 03: 0.10328; 04: 0.08165; 05: 0.13292; 06: 0.12247; 10: 0.35384; 11: 0.11372; 12: 0.13784; 13: 0.069402; 14: 0.063246
| Rel.Std.Dev. |  | 01: 1.249%; 02: 0.9604%; 03: 2.671%; 04: 2.168%; 05: 3.576%; 06: 3.101%; 10: 7.105%; 11: 2.967%; 12: 4.115%; 13: 1.861%; 14: 1.917%
| PDM3 |  | 01: 9.3%; 02: -3.569%; 03: 2.248%; 04: -0.3962%; 05: -1.718%; 06: 4.452%; 10: 31.69%; 11: 1.363%; 12: -11.41%; 13: -1.41%; 14: -12.74%
| Z-Score (Absolute) |  | 01: 2.06; 02: 0.67; 03: 0.56; 04: 0; 05: 0.28; 06: 1.03; 10: 6.82; 11: 0.37; 12: 2.34; 13: 0.22; 14: 2.62
| Analytical results for Ge in OREAS 554 (Certified Value 0.095 ppm) | [27 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 05 | 06 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | [10 empty] | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 | 1.4 | 1.1 | < 5 | < 0.1 | 0.07 | 0.07 | 0.11 | 0.1 | 0.08 | 0.08 | 0.2098 | < 0.1 | 1.5 | < 0.05 | 0.11 | 0.2 | [10 empty] | 1
| 1 | 2 | 1.5 | 0.9 | < 5 | < 0.1 | 0.06 | 0.11 | 0.13 | 0.1 | 0.09 | 0.08 | 0.18809 | < 0.1 | 1.5 | < 0.05 | 0.11 | 0.2 | [10 empty] | 0
| 1 | 3 | 1.4 | 0.9 | < 5 | < 0.1 | 0.06 | 0.08 | 0.11 | 0.09 | 0.09 | 0.07 | 0.17083 | < 0.1 | 1.6 | < 0.05 | 0.11 | 0.2 | [10 empty] | 16
| 1 | 4 | 1.5 | 1 | < 5 | < 0.1 | 0.06 | 0.08 | 0.12 | 0.1 | 0.13 | 0.07 | 0.195 | < 0.1 | 1.6 | < 0.05 | 0.11 | 0.3 | [10 empty] | 0.094857
| 1 | 5 | 1.4 | 1 | < 5 | < 0.1 | 0.12 | 0.08 | 0.13 | 0.1 | 0.12 | 0.08 | 0.20579 | < 0.1 | 1.6 | < 0.05 | 0.11 | 0.3 | [10 empty] | 101
| 1 | 6 | 1.4 | 1 | < 5 | < 0.1 | 0.11 | 0.16 | 0.11 | 0.1 | 0.08 | 0.07 | 0.18577 | 0.2 | 1.8 | < 0.05 | 0.11 | 0.3
| Mean |  | 01: 1.4333; 02: 0.98333; 05: 0.08; 06: 0.096667; 10: 0.11833; 11: 0.098333; 12: 0.098333; 13: 0.075; 14: 0.19255; 15: 0.2; 16: 1.6; 18: 0.11; 19: 0.25
| Median |  | 01: 1.4; 02: 1; 05: 0.065; 06: 0.08; 10: 0.115; 11: 0.1; 12: 0.09; 13: 0.075; 14: 0.19155; 15: 0.2; 16: 1.6; 18: 0.11; 19: 0.25
| Std Dev. |  | 01: 0.05164; 02: 0.075277; 05: 0.027568; 06: 0.033862; 10: 0.0098319; 11: 0.0040825; 12: 0.02137; 13: 0.0054772; 14: 0.014256; 16: 0.10954; 18: 0; 19: 0.054772
| Rel.Std.Dev. |  | 01: 3.603%; 02: 7.655%; 05: 34.46%; 06: 35.03%; 10: 8.309%; 11: 4.152%; 12: 21.73%; 13: 7.303%; 14: 7.404%; 16: 6.847%; 18: 0%; 19: 21.91%
| PDM3 |  | 01: 1411%; 02: 936.6%; 05: -15.66%; 06: 1.908%; 10: 24.75%; 11: 3.665%; 12: 3.665%; 13: -20.93%; 14: 103%; 15: 110.8%; 16: 1587%; 18: 15.96%; 19: 163.6%
| Z-Score (Absolute) |  | NA | NA | 74.77 | 1.48 | 0.54 | 0.03 | 0.65 | 0.03 | 0.03 | 0.7 | 2.96 | 0.7 | NA | 2.26 | 0.39 | NA
| Analytical results for Hf in OREAS 554 (Certified Value 1.88 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 1.91 | 1.97 | 1.6 | 1.8 | 1.7 | 2 | 1.6 | 1.9 | 1.85 | 1.8 | 1.9 | 1.6 | 1.8 | 1.7 | 1.83 | 2.2858 | 1.93 | 2.2565 | 1.7 | 1.87 | 2 | 2.18 | 1.8 |  |  |  | 1
| 1 | 2 | 1.84 | 2 | 1.6 | 1.6 | 1.7 | 2 | 2 | 2 | 1.81 | 2 | 2 | 1.7 | 1.8 | 1.7 | 1.82 | 2.2947 | 1.99 | 2.2733 | 1.6 | 1.88 | 1.9 | 2.22 | 1.8 |  |  |  | 0
| 1 | 3 | 1.82 | 1.96 | 1.8 | 1.8 | 1.7 | 2 | 2.1 | 2.1 | 1.84 | 1.8 | 1.9 | 1.7 | 1.8 | 1.7 | 1.78 | 2.3736 | 2.06 | 2.3591 | 1.6 | 1.76 | 1.9 | 2.13 | 1.8 |  |  |  | 16
| 1 | 4 | 1.88 | 2.02 | 1.6 | 1.6 | 1.7 | 2 | 2.2 | 2 | 1.93 | 1.8 | 2 | 1.7 | 1.8 | 1.6 | 1.74 | 2.2729 | 1.95 | 2.303 | 1.6 | 1.82 | 2 | 2.18 | 1.8 |  |  |  | 1.877
| 1 | 5 | 1.83 | 2.01 | 1.6 | 1.8 | 1.7 | 2 | 1.8 | 1.9 | 1.87 | 1.8 | 2.1 | 1.6 | 1.8 | 1.6 | 1.78 | 2.1767 | 2.02 | 2.2886 | 1.8 | 1.77 | 1.9 | 2.29 | 1.8 |  |  |  | 102
| 1 | 6 | 1.92 | 2 | 1.6 | 1.6 | 1.7 | 2 | 1.8 | 2 | 1.79 | 1.8 | 1.9 | 1.7 | 1.8 | 1.7 | 1.74 | 2.2053 | 2 | 2.1827 | 1.7 | 1.69 | 1.9 | 2.21 | 1.8
| Mean |  | 1.8667 | 1.9933 | 1.6333 | 1.7 | 1.7 | 2 | 1.9167 | 1.9833 | 1.8483 | 1.8333 | 1.9667 | 1.6667 | 1.8 | 1.6667 | 1.7817 | 2.2682 | 1.9917 | 2.2772 | 1.6667 | 1.7983 | 1.9333 | 2.2017 | 1.8
| Median |  | 1.86 | 2 | 1.6 | 1.7 | 1.7 | 2 | 1.9 | 2 | 1.845 | 1.8 | 1.95 | 1.7 | 1.8 | 1.7 | 1.78 | 2.2794 | 1.995 | 2.2809 | 1.65 | 1.795 | 1.9 | 2.195 | 1.8
| Std Dev. |  | 0.04274 | 0.023381 | 0.08165 | 0.10954 | 0 | 0 | 0.22286 | 0.075277 | 0.04916 | 0.08165 | 0.08165 | 0.05164 | 0 | 0.05164 | 0.038166 | 0.069951 | 0.047081 | 0.058082 | 0.08165 | 0.072503 | 0.05164 | 0.053448 | 0
| Rel.Std.Dev. |  | 2.29% | 1.173% | 4.999% | 6.444% | 0% | 0% | 11.63% | 3.795% | 2.66% | 4.454% | 4.152% | 3.098% | 0% | 3.098% | 2.142% | 3.084% | 2.364% | 2.551% | 4.899% | 4.032% | 2.671% | 2.428% | 0%
| PDM3 |  | -0.5498% | 6.199% | -12.98% | -9.429% | -9.429% | 6.554% | 2.114% | 5.666% | -1.527% | -2.326% | 4.778% | -11.21% | -4.102% | -11.21% | -5.078% | 20.84% | 6.11% | 21.32% | -11.21% | -4.19% | 3.002% | 17.3% | -4.102%
| Z-Score (Absolute) |  | 0.12 | 0.73 | 0.99 | 0.67 | 0.67 | NA | 0.36 | 0.68 | 0.04 | 0.04 | 0.6 | 0.83 | 0.2 | 0.83 | 0.28 | 2.05 | 0.72 | 2.09 | 0.83 | 0.2 | 0.44 | 1.73 | 0.2
Results from laboratory 6 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 554 (Indicative Value 0.053 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 01: Lab; 02: Lab |  |  |  | 1
| No | No | 01: 10; 02: 18 |  |  |  | ppm
|  |  | 01: 4A*OES/MS; 02: 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 01: 0.08; 02: NA |  |  |  | 1
| 1 | 2 | 01: 0.04; 02: NA |  |  |  | 0
| 1 | 3 | 01: 0.05; 02: NA |  |  |  | 16
| 1 | 4 | 01: 0.04; 02: NA |  |  |  | 0.053333
| 1 | 5 | 01: 0.03; 02: NA |  |  |  | 25
| 1 | 6 | 01: 0.08; 02: NA
| Mean |  | 01: 0.053333
| Median |  | 01: 0.045
| Std Dev. |  | 01: 0.021602
| Rel.Std.Dev. |  | 01: 40.5%
| PDM3 |  | 01: 6.661e-14%
Z-Score (Absolute) | bar: 01 to 02
| Analytical results for Ho in OREAS 554 (Certified Value 0.52 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 19; 11: 22; 13: 24; 14: 29; 15: 30 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*MS |  |  |  | 2
2
| 1 | 1 | 01: 0.48; 02: 0.48; 03: 0.56; 04: 0.54; 05: 0.45; 06: 0.6; 09: 0.61; 10: 0.57895; 11: 0.6447; 13: 0.4; 14: 0.46; 15: 0.41 |  |  |  | 1
| 1 | 2 | 01: 0.48; 02: 0.51; 03: 0.54; 04: 0.54; 05: 0.4; 06: 0.6; 09: 0.51; 10: 0.60778; 11: 0.6474; 13: 0.4; 14: 0.43; 15: 0.41 |  |  |  | 0
| 1 | 3 | 01: 0.46; 02: 0.51; 03: 0.54; 04: 0.54; 05: 0.4; 06: 0.6; 09: 0.54; 10: 0.62002; 11: 0.65721; 13: 0.4; 14: 0.47; 15: 0.41 |  |  |  | 16
| 1 | 4 | 01: 0.45; 02: 0.48; 03: 0.54; 04: 0.56; 05: 0.4; 06: 0.6; 09: 0.53; 10: 0.60096; 11: 0.66039; 13: 0.5; 14: 0.45; 15: 0.48 |  |  |  | 0.52471
| 1 | 5 | 01: 0.46; 02: 0.48; 03: 0.54; 04: 0.54; 05: 0.4; 06: 0.6; 09: 0.64; 10: 0.57391; 11: 0.67199; 13: 0.5; 14: 0.45; 15: 0.4 |  |  |  | 103
| 1 | 6 | 01: 0.46; 02: 0.5; 03: 0.56; 04: 0.54; 05: 0.4; 06: 0.6; 09: 0.52; 10: 0.58563; 11: 0.6855; 13: 0.5; 14: 0.44; 15: 0.42
| Mean |  | 01: 0.465; 02: 0.49333; 03: 0.54667; 04: 0.54333; 05: 0.40833; 06: 0.6; 09: 0.55833; 10: 0.59454; 11: 0.6612; 13: 0.45; 14: 0.45; 15: 0.42167
| Median |  | 01: 0.46; 02: 0.49; 03: 0.54; 04: 0.54; 05: 0.4; 06: 0.6; 09: 0.535; 10: 0.5933; 11: 0.6588; 13: 0.45; 14: 0.45; 15: 0.41
| Std Dev. |  | 01: 0.012247; 02: 0.015055; 03: 0.010328; 04: 0.008165; 05: 0.020412; 06: 0; 09: 0.053448; 10: 0.017965; 11: 0.015402; 13: 0.054772; 14: 0.014142; 15: 0.029269
| Rel.Std.Dev. |  | 01: 2.634%; 02: 3.052%; 03: 1.889%; 04: 1.503%; 05: 4.999%; 06: 0%; 09: 9.573%; 10: 3.022%; 11: 2.329%; 13: 12.17%; 14: 3.143%; 15: 6.941%
| PDM3 |  | 01: -11.38%; 02: -5.98%; 03: 4.184%; 04: 3.549%; 05: -22.18%; 06: 14.35%; 09: 6.408%; 10: 13.31%; 11: 26.01%; 13: -14.24%; 14: -14.24%; 15: -19.64%
| Z-Score (Absolute) |  | 01: 1.03; 02: 0.66; 03: 0.04; 04: 0; 05: NA; 06: NA; 09: 0.2; 10: 0.67; 11: 1.55; 13: NA; 14: 1.23; 15: 1.6
Results from laboratories 6, 10 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 554 (Certified Value 0.1 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 28 | 29 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.11 | 0.1 | 0.12 | 0.1 | 0.1 | 0.1 | < 0.1 | 0.12 | 17 | 0.104 | 0.114 | 0.104 | 0.103 | 0.1 | 0.12 | < 0.5 | 0.1 | 0.10003 | 0.11 | 0.09 | 0.109 | 0.11 |  |  |  |  | 1
| 1 | 2 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | < 0.1 | 0.12 | 18 | 0.108 | 0.118 | 0.1 | 0.117 | 0.105 | 0.12 | < 0.5 | 0.11 | 0.10006 | 0.1 | 0.1 | 0.103 | 0.11 |  |  |  |  | 0
| 1 | 3 | 0.11 | 0.11 | 0.1 | 0.1 | 0.1 | 0.1 | < 0.1 | 0.11 | 18 | 0.105 | 0.111 | 0.112 | 0.109 | 0.098 | 0.12 | < 0.5 | 0.1 | 0.09602 | 0.11 | 0.1 | 0.109 | 0.11 |  |  |  |  | 16
| 1 | 4 | 0.11 | 0.11 | 0.12 | 0.1 | 0.11 | 0.1 | 0.1 | 0.11 | 17 | 0.104 | 0.112 | 0.111 | 0.108 | 0.12 | 0.12 | < 0.5 | 0.1 | 0.09897 | 0.1 | 0.1 | 0.105 | 0.11 |  |  |  |  | 0.10494
| 1 | 5 | 0.1 | 0.1 | 0.1 | 0.1 | 0.11 | < 0.1 | < 0.1 | 0.12 | 17 | 0.101 | 0.106 | 0.108 | 0.105 | 0.093 | 0.12 | < 0.5 | 0.11 | 0.10251 | 0.1 | 0.1 | 0.105 | 0.11 |  |  |  |  | 104
| 1 | 6 | 0.1 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | 0.1 | 0.11 | 17 | 0.097 | 0.102 | 0.104 | 0.114 | 0.099 | 0.12 | < 0.5 | 0.11 | 0.10128 | 0.1 | 0.11 | 0.103 | 0.11
| Mean |  | 0.105 | 0.10333 | 0.10667 | 0.1 | 0.10167 | 0.1 | 0.1 | 0.115 | 17.3333 | 0.10317 | 0.1105 | 0.1065 | 0.10933 | 0.1025 | 0.12 |  | 0.105 | 0.099812 | 0.10333 | 0.1 | 0.10567 | 0.11
| Median |  | 0.105 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.1 | 0.115 | 17 | 0.104 | 0.1115 | 0.106 | 0.1085 | 0.0995 | 0.12 |  | 0.105 | 0.10004 | 0.1 | 0.1 | 0.105 | 0.11
| Std Dev. |  | 0.0054772 | 0.005164 | 0.010328 | 1.5202e-17 | 0.0075277 | 0 | 0 | 0.0054772 | 0.5164 | 0.0037639 | 0.0057184 | 0.0046368 | 0.0053166 | 0.0093968 | 0 |  | 0.0054772 | 0.0022195 | 0.005164 | 0.0063246 | 0.0027325 | 0
| Rel.Std.Dev. |  | 5.216% | 4.997% | 9.682% | 1.52e-14% | 7.404% | 0% | 0% | 4.763% | 2.979% | 3.648% | 5.175% | 4.354% | 4.863% | 9.168% | 0% |  | 5.216% | 2.224% | 4.997% | 6.325% | 2.586% | 0%
| PDM3 |  | 0.05727% | -1.531% | 1.645% | -4.707% | -3.119% | -4.707% | -4.707% | 9.587% | 16417% | -1.69% | 5.298% | 1.487% | 4.187% | -2.325% | 14.35% |  | 0.05727% | -4.887% | -1.531% | -4.707% | 0.6926% | 4.822%
| Z-Score (Absolute) |  | 0 | 0.34 | 0.34 | NA | 0.67 | NA | 7.75 | 2.02 | NA | 0.37 | 1.11 | 0.3 | 0.88 | 0.51 | 3.03 | 29.33 | 0 | 1.05 | 0.34 | 1.01 | 0.13 | 1.01
Results from laboratories 4, 6 & 12 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for K in OREAS 554 (Certified Value 1.71 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 1.77 | 1.6297 | 1.72 | 1.72 | 1.64 | 1.76 | 1.67 | 1.98 | 1.74 | 1.79 | 1.68 | 1.72 | 1.61 | 1.7 | 1.71 | 1.7331 | 1.878 | 1.79 | 1.76 | 1.56 | 1.6625 | 1.7 | 1.71 | 1.6354 |  |  | 1
| 1 | 2 | 1.79 | 1.6227 | 1.71 | 1.68 | 1.629 | 1.73 | 1.69 | 1.9 | 1.69 | 1.78 | 1.74 | 1.72 | 1.67 | 1.69 | 1.7 | 1.7317 | 1.9056 | 1.83 | 1.75 | 1.61 | 1.672 | 1.71 | 1.68 | 1.6188 |  |  | 0
| 1 | 3 | 1.79 | 1.6142 | 1.72 | 1.7 | 1.635 | 1.72 | 1.69 | 2.03 | 1.76 | 1.79 | 1.71 | 1.69 | 1.65 | 1.65 | 1.66 | 1.7724 | 1.45 | 1.83 | 1.81 | 1.55 | 1.672 | 1.69 | 1.68 | 1.6437 |  |  | 16
| 1 | 4 | 1.79 | 1.6313 | 1.69 | 1.68 | 1.632 | 1.76 | 1.67 | 1.63 | 1.72 | 1.75 | 1.69 | 1.69 | 1.65 | 1.69 | 1.62 | 1.7265 | 1.9116 | 1.78 | 1.77 | 1.61 | 1.729 | 1.72 | 1.69 | 1.6188 |  |  | 1.7074
| 1 | 5 | 1.79 | 1.6358 | 1.72 | 1.68 | 1.645 | 1.75 | 1.67 | 1.9 | 1.74 | 1.77 | 1.66 | 1.7 | 1.67 | 1.67 | 1.69 | 1.7661 | 1.806 | 1.83 | 1.81 | 1.56 | 1.653 | 1.71 | 1.68 | 1.6437 |  |  | 105
| 1 | 6 | 1.8 | 1.6585 | 1.73 | 1.69 | 1.649 | 1.74 | 1.66 | 2.01 | 1.72 | 1.77 | 1.65 | 1.62 | 1.67 | 1.73 | 1.65 | 1.7391 | 1.8216 | 1.82 | 1.76 | 1.55 | 1.7385 | 1.72 | 1.69 | 1.6354
| Mean |  | 1.7883 | 1.632 | 1.715 | 1.6917 | 1.6383 | 1.7433 | 1.675 | 1.9083 | 1.7283 | 1.775 | 1.6883 | 1.69 | 1.6533 | 1.6883 | 1.6717 | 1.7448 | 1.7955 | 1.8133 | 1.7767 | 1.5733 | 1.6878 | 1.7083 | 1.6883 | 1.6326
| Median |  | 1.79 | 1.6305 | 1.72 | 1.685 | 1.6375 | 1.745 | 1.67 | 1.94 | 1.73 | 1.775 | 1.685 | 1.695 | 1.66 | 1.69 | 1.675 | 1.7361 | 1.8498 | 1.825 | 1.765 | 1.56 | 1.672 | 1.71 | 1.685 | 1.6354
| Std Dev. |  | 0.0098319 | 0.015005 | 0.013784 | 0.016021 | 0.0077374 | 0.01633 | 0.012247 | 0.14689 | 0.024014 | 0.015166 | 0.033116 | 0.036878 | 0.023381 | 0.027142 | 0.034303 | 0.019449 | 0.17468 | 0.022509 | 0.026583 | 0.028752 | 0.036382 | 0.01169 | 0.01169 | 0.01134
| Rel.Std.Dev. |  | 0.5498% | 0.9194% | 0.8037% | 0.947% | 0.4723% | 0.9367% | 0.7312% | 7.697% | 1.389% | 0.8544% | 1.961% | 2.182% | 1.414% | 1.608% | 2.052% | 1.115% | 9.729% | 1.241% | 1.496% | 1.827% | 2.156% | 0.6843% | 0.6924% | 0.6946%
| PDM3 |  | 4.739% | -4.416% | 0.4436% | -0.9229% | -4.047% | 2.103% | -1.899% | 11.77% | 1.225% | 3.958% | -1.118% | -1.021% | -3.168% | -1.118% | -2.094% | 2.19% | 5.156% | 6.203% | 4.055% | -7.853% | -1.147% | 0.05319% | -1.118% | -4.38%
| Z-Score (Absolute) |  | 1.46 | 0.88 | 0.36 | 0.01 | 0.79 | 0.79 | 0.24 | 3.26 | 0.56 | 1.26 | 0.04 | 0.01 | 0.56 | 0.04 | 0.29 | 0.81 | 1.57 | 1.84 | 1.29 | 1.76 | 0.04 | 0.26 | 0.04 | 0.87
| Analytical results for La in OREAS 554 (Certified Value 30.6 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 24 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 27.9 | 31.1 | 32.1 | 30.6 | 30 | 28.9 | 32.4 | 25.2 | 24 | 32.5 | 34.4 | 31.1 | 33.2 | 29.8 | 30.05 | 37.8841 | 31.2 | 33.3341 | 29.2 | 31.6 | 26.5 | 28 |  |  |  |  | 1
| 1 | 2 | 27.6 | 31.5 | 31.6 | 30.7 | 29 | 28.6 | 32.9 | 24.8 | 23 | 32.6 | 36.1 | 31.9 | 33.8 | 30.7 | 30.37 | 38.2481 | 33.2 | 33.4567 | 29.9 | 31.6 | 27.1 | 31 |  |  |  |  | 0
| 1 | 3 | 27.8 | 31.8 | 32 | 30.1 | 28.5 | 28.9 | 34.4 | 24.7 | 23 | 32.6 | 34.4 | 31.5 | 32.9 | 29.9 | 30.52 | 38.4858 | 33.7 | 33.9055 | 29.6 | 31.8 | 27.5 | 29 |  |  |  |  | 16
| 1 | 4 | 27.4 | 31.47 | 30.8 | 30.3 | 30.5 | 27.6 | 32.8 | 25.8 | 24 | 31.5 | 34.3 | 30.7 | 33.4 | 30.1 | 29.93 | 37.7504 | 32.1 | 33.5563 | 30.5 | 31.5 | 26.4 | 27 |  |  |  |  | 30.5763
| 1 | 5 | 27.3 | 31.12 | 31.6 | 30.5 | 29.5 | 28.8 | 33.7 | 24.4 | 23 | 31.6 | 34.6 | 30.6 | 33.5 | 29.8 | 30.64 | 37.6003 | 32.1 | 33.4146 | 28.8 | 31.2 | 26.5 | 29 |  |  |  |  | 106
| 1 | 6 | 27.7 | 31.18 | 32.3 | 31.9 | 28.5 | 28.7 | 33.3 | 24.1 | 24 | 31.9 | 33.6 | 32.6 | 33.2 | 29.4 | 30.62 | 38.1913 | 31.3 | 32.9083 | 30.5 | 31.9 | 27.3 | 29
| Mean |  | 27.6167 | 31.3617 | 31.7333 | 30.6833 | 29.3333 | 28.5833 | 33.25 | 24.8333 | 23.5 | 32.1167 | 34.5667 | 31.4 | 33.3333 | 29.95 | 30.355 | 38.0267 | 32.2667 | 33.4292 | 29.75 | 31.6 | 26.8833 | 28.8333
| Median |  | 27.65 | 31.325 | 31.8 | 30.55 | 29.25 | 28.75 | 33.1 | 24.75 | 23.5 | 32.2 | 34.4 | 31.3 | 33.3 | 29.85 | 30.445 | 38.0377 | 32.1 | 33.4356 | 29.75 | 31.6 | 26.8 | 29
| Std Dev. |  | 0.23166 | 0.27672 | 0.53541 | 0.63377 | 0.8165 | 0.49565 | 0.71764 | 0.60222 | 0.54772 | 0.51153 | 0.82624 | 0.7642 | 0.30768 | 0.43243 | 0.30085 | 0.33627 | 1.0053 | 0.32386 | 0.6892 | 0.24495 | 0.47504 | 1.3292
| Rel.Std.Dev. |  | 0.8388% | 0.8824% | 1.687% | 2.066% | 2.784% | 1.734% | 2.158% | 2.425% | 2.331% | 1.593% | 2.39% | 2.434% | 0.923% | 1.444% | 0.9911% | 0.8843% | 3.116% | 0.9688% | 2.317% | 0.7752% | 1.767% | 4.61%
| PDM3 |  | -9.679% | 2.569% | 3.784% | 0.3501% | -4.065% | -6.518% | 8.744% | -18.78% | -23.14% | 5.038% | 13.05% | 2.694% | 9.017% | -2.048% | -0.7237% | 24.37% | 5.528% | 9.331% | -2.702% | 3.348% | -12.08% | -5.7%
| Z-Score (Absolute) |  | 1.18 | 0.12 | 0.25 | 0.12 | 0.59 | 0.85 | 0.77 | 2.15 | 2.62 | 0.38 | 1.23 | 0.13 | 0.8 | 0.37 | 0.23 | 2.44 | 0.43 | 0.84 | 0.44 | 0.2 | 1.44 | 0.76
| Analytical results for Li in OREAS 554 (Certified Value 33.9 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 35 | 33.7 | 34 | 40 | 37.5 | 36 | 35.4 | 33.7 | 35.1 | 41 | 34.6 | 33.8 | 33.3 | 34.2 | 31.2 | 36.5 | 35 | 41 | 33.0168 | 31 | 33.5 | 32.4 | 34 | 32 |  |  | 1
| 1 | 2 | 35 | 32.9 | 33 | 40 | 36 | 36 | 34 | 33.3 | 34.7 | 42 | 34.4 | 33.2 | 34 | 34.4 | 32 | 37 | 35.4 | 41 | 33.2402 | 31.7 | 34 | 32.3 | 32 | 31 |  |  | 0
| 1 | 3 | 35 | 33.3 | 33 | 30 | 35.6 | 36 | 35 | 35.9 | 35.6 | 42 | 33.9 | 31.8 | 33.3 | 33.9 | 31.1 | 34.8 | 35.4 | 43 | 32.0608 | 31.5 | 34 | 32.5 | 34 | 32 |  |  | 16
| 1 | 4 | 36 | 33.6 | 32 | 40 | 36.8 | 36 | 34.8 | 33.6 | 36.2 | 42 | 33.1 | 32.8 | 33.3 | 34.9 | 31.5 | 33.9 | 34.9 | 43 | 32.5906 | 31.4 | 35.2 | 32 | 33 | 31 |  |  | 33.87
| 1 | 5 | 36 | 33 | 32.5 | 30 | 37.2 | 36 | 34.5 | 34 | 36.7 | 43 | 33.2 | 32.1 | 33.3 | 34.1 | 31.8 | 34.7 | 35.8 | 41 | 32.5646 | 31.3 | 34.1 | 32 | 32 | 33 |  |  | 107
| 1 | 6 | 36 | 33.4 | 33 | 30 | 37.6 | 36 | 32.1 | 34.9 | 36 | 42 | 33.7 | 32.4 | 33.8 | 34.7 | 30.5 | 36.5 | 35.6 | 40 | 32.5046 | 30.6 | 35.8 | 32.3 | 34 | 31
| Mean |  | 35.5 | 33.3167 | 32.9167 | 35 | 36.7833 | 36 | 34.3 | 34.2333 | 35.7167 | 42 | 33.8167 | 32.6833 | 33.5 | 34.3667 | 31.35 | 35.5667 | 35.35 | 41.5 | 32.6629 | 31.25 | 34.4333 | 32.25 | 33.1667 | 31.6667
| Median |  | 35.5 | 33.35 | 33 | 35 | 37 | 36 | 34.65 | 33.85 | 35.8 | 42 | 33.8 | 32.6 | 33.3 | 34.3 | 31.35 | 35.65 | 35.4 | 41 | 32.5776 | 31.35 | 34.05 | 32.3 | 33.5 | 31.5
| Std Dev. |  | 0.54772 | 0.31885 | 0.66458 | 5.4772 | 0.82077 | 0 | 1.1764 | 0.98319 | 0.73598 | 0.63246 | 0.61128 | 0.73869 | 0.31623 | 0.37771 | 0.53944 | 1.258 | 0.34496 | 1.2247 | 0.41497 | 0.3937 | 0.87331 | 0.20736 | 0.98319 | 0.8165
| Rel.Std.Dev. |  | 1.543% | 0.957% | 2.019% | 15.65% | 2.231% | 0% | 3.43% | 2.872% | 2.061% | 1.506% | 1.808% | 2.26% | 0.944% | 1.099% | 1.721% | 3.537% | 0.9759% | 2.951% | 1.27% | 1.26% | 2.536% | 0.643% | 2.964% | 2.578%
| PDM3 |  | 4.813% | -1.634% | -2.815% | 3.336% | 8.602% | 6.289% | 1.27% | 1.073% | 5.452% | 24% | -0.1574% | -3.504% | -1.092% | 1.466% | -7.44% | 5.009% | 4.37% | 22.53% | -3.564% | -7.735% | 1.663% | -4.783% | -2.077% | -6.505%
| Z-Score (Absolute) |  | 0.64 | 0.46 | 0.67 | NA | 1.29 | 0.89 | 0.03 | 0 | 0.75 | 3.93 | 0.21 | 0.78 | 0.37 | 0.07 | 1.46 | 0.67 | 0.56 | 3.67 | 0.79 | 1.51 | 0.1 | 1 | 0.54 | 1.3
Results from laboratory 4 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 554 (Certified Value 0.21 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 21; 11: 22; 12: 24; 13: 29 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 01: 0.21; 02: 0.21; 03: 0.24; 04: 0.22; 05: < 0.1; 06: 0.2; 09: 0.22; 10: 0.23; 11: 0.26796; 12: 0.2; 13: 0.18 |  |  | 1
| 1 | 2 | 01: 0.21; 02: 0.19; 03: 0.24; 04: 0.2; 05: < 0.1; 06: 0.2; 09: 0.21; 10: 0.23; 11: 0.26343; 12: 0.2; 13: 0.19 |  |  | 0
| 1 | 3 | 01: 0.21; 02: 0.18; 03: 0.22; 04: 0.22; 05: < 0.1; 06: 0.3; 09: 0.22; 10: 0.23; 11: 0.2701; 12: 0.2; 13: 0.19 |  |  | 16
| 1 | 4 | 01: 0.21; 02: 0.2; 03: 0.22; 04: 0.22; 05: < 0.1; 06: 0.2; 09: 0.22; 10: 0.22; 11: 0.27258; 12: 0.2; 13: 0.19 |  |  | 0.21143
| 1 | 5 | 01: 0.2; 02: 0.2; 03: 0.22; 04: 0.22; 05: 0.1; 06: 0.2; 09: 0.21; 10: 0.23; 11: 0.2645; 12: 0.2; 13: 0.19 |  |  | 108
| 1 | 6 | 01: 0.21; 02: 0.2; 03: 0.22; 04: 0.22; 05: < 0.1; 06: 0.2; 09: 0.22; 10: 0.22; 11: 0.27601; 12: 0.2; 13: 0.19
| Mean |  | 01: 0.20833; 02: 0.19667; 03: 0.22667; 04: 0.21667; 05: 0.1; 06: 0.21667; 09: 0.21667; 10: 0.22667; 11: 0.2691; 12: 0.2; 13: 0.18833
| Median |  | 01: 0.21; 02: 0.2; 03: 0.22; 04: 0.22; 05: 0.1; 06: 0.2; 09: 0.22; 10: 0.23; 11: 0.26903; 12: 0.2; 13: 0.19
| Std Dev. |  | 01: 0.0040825; 02: 0.010328; 03: 0.010328; 04: 0.008165; 06: 0.040825; 09: 0.005164; 10: 0.005164; 11: 0.0048056; 12: 3.0405e-17; 13: 0.0040825
| Rel.Std.Dev. |  | 01: 1.96%; 02: 5.252%; 03: 4.556%; 04: 3.768%; 06: 18.84%; 09: 2.383%; 10: 2.278%; 11: 1.786%; 12: 1.52e-14%; 13: 2.168%
| PDM3 |  | 01: -1.464%; 02: -6.982%; 03: 7.207%; 04: 2.477%; 05: -52.7%; 06: 2.477%; 09: 2.477%; 10: 7.207%; 11: 27.28%; 12: -5.405%; 13: -10.92%
| Z-Score (Absolute) |  | 01: 0.56; 02: 1.35; 03: 0.67; 04: 0; 05: 10.68; 06: NA; 09: 0; 10: 0.67; 11: 3.54; 12: NA; 13: 1.91
Results from laboratories 10 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 554 (Certified Value 3.14 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 3.18 | 3.0452 | 3.16 | 3.06 | 3.097 | 3.185 | 3.1 | 3.29 | 3.21 | 3.11 | 3.18 | 3.28 | 3 | 3.09 | 3.14 | 3.1579 | 3.153 | 3.1 | 3.12 | 3.43 | 3.0456 | 3.27 | 2.89 | 3.1721 |  |  | 1
| 1 | 2 | 3.22 | 3.09 | 3.1 | 3.03 | 3.086 | 3.1525 | 3.13 | 3.19 | 3.18 | 3.15 | 3.25 | 3.2 | 3.15 | 3.06 | 3.11 | 3.1212 | 3.075 | 3.19 | 3.17 | 3.33 | 3.0644 | 3.26 | 2.9 | 3.1962 |  |  | 0
| 1 | 3 | 3.22 | 3.0195 | 3.16 | 3.07 | 3.144 | 3.1665 | 3.13 | 3.43 | 3.31 | 3.09 | 3.23 | 3.13 | 3.07 | 3 | 3.07 | 3.2232 | 3.093 | 3.17 | 3.22 | 3.26 | 3.0832 | 3.25 | 3 | 3.2143 |  |  | 16
| 1 | 4 | 3.24 | 3.0807 | 3.07 | 3.05 | 3.156 | 3.2045 | 3.09 | 3.33 | 3.23 | 3.16 | 3.14 | 3.14 | 3.04 | 3.08 | 2.99 | 3.1189 | 3.13 | 3.11 | 3.16 | 3.23 | 3.1772 | 3.3 | 2.98 | 3.154 |  |  | 3.1442
| 1 | 5 | 3.25 | 3.0675 | 3.11 | 3 | 3.193 | 3.173 | 3.08 | 3.3 | 3.4 | 3.13 | 3.06 | 3.19 | 3.05 | 3.03 | 3.1 | 3.1911 | 3.02 | 3.19 | 3.21 | 3.33 | 3.0268 | 3.27 | 3.05 | 3.2083 |  |  | 109
| 1 | 6 | 3.24 | 3.0595 | 3.13 | 3.02 | 3.227 | 3.1745 | 3.07 | 3.32 | 3.2 | 3.1 | 3.09 | 3.15 | 3.1 | 3.13 | 3.01 | 3.1251 | 3.018 | 3.16 | 3.12 | 3.34 | 3.2712 | 3.19 | 2.9 | 3.1842
| Mean |  | 3.225 | 3.0604 | 3.1217 | 3.0383 | 3.1505 | 3.176 | 3.1 | 3.31 | 3.255 | 3.1233 | 3.1583 | 3.1817 | 3.0683 | 3.065 | 3.07 | 3.1562 | 3.0815 | 3.1533 | 3.1667 | 3.32 | 3.1114 | 3.2567 | 2.9533 | 3.1882
| Median |  | 3.23 | 3.0635 | 3.12 | 3.04 | 3.15 | 3.1738 | 3.095 | 3.31 | 3.22 | 3.12 | 3.16 | 3.17 | 3.06 | 3.07 | 3.085 | 3.1415 | 3.084 | 3.165 | 3.165 | 3.33 | 3.0738 | 3.265 | 2.94 | 3.1902
| Std Dev. |  | 0.0251 | 0.025469 | 0.035449 | 0.026394 | 0.05432 | 0.01759 | 0.025298 | 0.077201 | 0.084083 | 0.028048 | 0.075741 | 0.055648 | 0.051929 | 0.045935 | 0.058992 | 0.043112 | 0.055594 | 0.039328 | 0.04274 | 0.069857 | 0.094188 | 0.036697 | 0.066232 | 0.022777
| Rel.Std.Dev. |  | 0.7783% | 0.8322% | 1.136% | 0.8687% | 1.724% | 0.5538% | 0.8161% | 2.332% | 2.583% | 0.898% | 2.398% | 1.749% | 1.692% | 1.499% | 1.922% | 1.366% | 1.804% | 1.247% | 1.35% | 2.104% | 3.027% | 1.127% | 2.243% | 0.7144%
| PDM3 |  | 2.568% | -2.667% | -0.7181% | -3.368% | 0.1989% | 1.01% | -1.407% | 5.272% | 3.522% | -0.6651% | 0.448% | 1.19% | -2.414% | -2.52% | -2.361% | 0.3813% | -1.996% | 0.289% | 0.7131% | 5.59% | -1.045% | 3.575% | -6.072% | 1.397%
| Z-Score (Absolute) |  | 0.81 | 1.01 | 0.33 | 1.25 | 0.02 | 0.27 | 0.57 | 1.74 | 1.14 | 0.32 | 0.07 | 0.33 | 0.92 | 0.96 | 0.9 | 0.05 | 0.78 | 0.02 | 0.16 | 1.85 | 0.45 | 1.15 | 2.19 | 0.4
| Analytical results for Mn in OREAS 554 (Certified Value 0.108 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.1019 | 0.1108 | 0.112 | 0.114 | 0.1092 | 0.1145 | 0.1104 | 0.105 | 0.106 | 0.0955 | 0.11 | 0.113 | 0.106 | 0.106 | 0.103 | 0.1144 | 0.1019 | 0.108 | 0.11 | 0.11 | 0.1071 | 0.1094 | 0.1008 | 0.11 |  |  | 1
| 1 | 2 | 0.1038 | 0.1128 | 0.111 | 0.111 | 0.1082 | 0.1135 | 0.113 | 0.102 | 0.104 | 0.0948 | 0.113 | 0.11 | 0.112 | 0.104 | 0.106 | 0.1119 | 0.1065 | 0.113 | 0.11 | 0.104 | 0.1071 | 0.1076 | 0.0994 | 0.11 |  |  | 0
| 1 | 3 | 0.1043 | 0.1132 | 0.115 | 0.11 | 0.1076 | 0.114 | 0.1127 | 0.111 | 0.106 | 0.0952 | 0.112 | 0.109 | 0.11 | 0.103 | 0.107 | 0.1169 | 0.1077 | 0.11 | 0.11 | 0.103 | 0.1089 | 0.1064 | 0.0998 | 0.1123 |  |  | 16
| 1 | 4 | 0.1061 | 0.1142 | 0.111 | 0.112 | 0.1074 | 0.113 | 0.1096 | 0.101 | 0.108 | 0.094 | 0.107 | 0.106 | 0.106 | 0.105 | 0.0991 | 0.112 | 0.1043 | 0.109 | 0.11 | 0.103 | 0.1107 | 0.111 | 0.1018 | 0.1123 |  |  | 0.10817
| 1 | 5 | 0.1059 | 0.1115 | 0.113 | 0.111 | 0.1079 | 0.113 | 0.1085 | 0.103 | 0.106 | 0.0957 | 0.106 | 0.109 | 0.107 | 0.102 | 0.103 | 0.1157 | 0.09979 | 0.111 | 0.11 | 0.106 | 0.1053 | 0.1081 | 0.1023 | 0.1107 |  |  | 110
| 1 | 6 | 0.1072 | 0.1121 | 0.115 | 0.111 | 0.109 | 0.1155 | 0.109 | 0.101 | 0.107 | 0.0946 | 0.108 | 0.107 | 0.108 | 0.108 | 0.1 | 0.1118 | 0.1004 | 0.109 | 0.11 | 0.105 | 0.1107 | 0.1039 | 0.1026 | 0.1131
| Mean |  | 0.10487 | 0.11243 | 0.11283 | 0.1115 | 0.10822 | 0.11392 | 0.11053 | 0.10383 | 0.10617 | 0.094967 | 0.10933 | 0.109 | 0.10817 | 0.10467 | 0.10302 | 0.11378 | 0.10343 | 0.11 | 0.11 | 0.10517 | 0.1083 | 0.10773 | 0.10112 | 0.1114
| Median |  | 0.1051 | 0.11245 | 0.1125 | 0.111 | 0.10805 | 0.11375 | 0.11 | 0.1025 | 0.106 | 0.095 | 0.109 | 0.109 | 0.1075 | 0.1045 | 0.103 | 0.1132 | 0.1031 | 0.1095 | 0.11 | 0.1045 | 0.108 | 0.10785 | 0.1013 | 0.1115
| Std Dev. |  | 0.0019128 | 0.0012242 | 0.0018348 | 0.0013784 | 0.00073869 | 0.0009704 | 0.0019054 | 0.0038166 | 0.0013292 | 0.00062823 | 0.0028048 | 0.0024495 | 0.0024014 | 0.0021602 | 0.0031371 | 0.0022104 | 0.0032615 | 0.0017889 | 0 | 0.0026394 | 0.0021799 | 0.0024525 | 0.0013303 | 0.0013357
| Rel.Std.Dev. |  | 1.824% | 1.089% | 1.626% | 1.236% | 0.6826% | 0.8518% | 1.724% | 3.676% | 1.252% | 0.6615% | 2.565% | 2.247% | 2.22% | 2.064% | 3.045% | 1.943% | 3.153% | 1.626% | 0% | 2.51% | 2.013% | 2.276% | 1.316% | 1.199%
| PDM3 |  | -3.056% | 3.939% | 4.308% | 3.076% | 0.04055% | 5.31% | 2.182% | -4.012% | -1.855% | -12.21% | 1.073% | 0.7647% | -0.00567% | -3.241% | -4.767% | 5.187% | -4.383% | 1.689% | 1.689% | -2.779% | 0.1176% | -0.4063% | -6.523% | 2.983%
| Z-Score (Absolute) |  | 0.72 | 0.88 | 0.97 | 0.69 | 0.01 | 1.2 | 0.48 | 0.94 | 0.44 | 2.81 | 0.23 | 0.16 | 0.02 | 0.76 | 1.11 | 1.17 | 1.02 | 0.37 | 0.37 | 0.65 | 0.01 | 0.11 | 1.51 | 0.66
| Analytical results for Mo in OREAS 554 (Certified Value 3.54 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  | 2
3
| 1 | 1 | 3.89 | 3.6 | 3.5 | 3.5 | 3.3 | 4 | 3.3 | 3.75 | 3.64 | 3 | 3.65 | 3.51 | 3.63 | 3.41 | 3.69 | 3.9 | 3.9698 | 3.12 | 3.57 | 3.3712 | 3.3 | 3.56 | 3.53 | 3.42 | 2.91 |  | 1
| 1 | 2 | 3.72 | 3.8 | 3.5 | 3.5 | 3.3 | 3 | 3.4 | 3.4 | 3.65 | 3 | 3.66 | 3.52 | 3.88 | 3.45 | 3.89 | 3.9 | 4.3475 | 3.12 | 3.7 | 3.4289 | 3.1 | 3.48 | 3.52 | 3.37 | 3.08 |  | 0
| 1 | 3 | 3.79 | 3.5 | 3.5 | 3.5 | 3.2 | 3 | 3.4 | 3.97 | 3.55 | 3 | 3.57 | 3.28 | 3.67 | 3.45 | 3.71 | 3.9 | 4 | 3.3 | 3.74 | 3.3692 | 3.3 | 3.44 | 3.5 | 3.49 | 3.05 |  | 16
| 1 | 4 | 3.74 | 3.8 | 3.5 | 3.5 | 3.3 | 3 | 3.3 | 3.53 | 3.7 | 3 | 3.51 | 3.61 | 3.72 | 3.53 | 3.8 | 3.7 | 4.0323 | 3.17 | 6.43 | 3.4349 | 3.3 | 3.36 | 3.58 | 3.4 | 3.28 |  | 3.5386
| 1 | 5 | 3.82 | 3.6 | 3.5 | 3.5 | 3.3 | 4 | 3.4 | 3.52 | 3.58 | 3 | 3.57 | 3.44 | 3.78 | 3.57 | 3.94 | 3.9 | 3.9176 | 3.14 | 3.81 | 3.5358 | 3.3 | 3.71 | 3.57 | 3.51 | 3.26 |  | 111
| 1 | 6 | 3.74 | 3.7 | 3.5 | 4 | 3.3 | 4 | 3.5 | 3.49 | 3.57 | 3 | 3.47 | 3.66 | 3.72 | 3.55 | 3.81 | 3.8 | 3.8955 | 3.25 | 3.71 | 3.2404 | 3.4 | 3.73 | 3.57 | 3.32 | 3.13
| Mean |  | 3.7833 | 3.6667 | 3.5 | 3.5833 | 3.2833 | 3.5 | 3.3833 | 3.61 | 3.615 | 3 | 3.5717 | 3.5033 | 3.7333 | 3.4933 | 3.8067 | 3.85 | 4.0271 | 3.1833 | 4.16 | 3.3967 | 3.2833 | 3.5467 | 3.545 | 3.4183 | 3.1183
| Median |  | 3.765 | 3.65 | 3.5 | 3.5 | 3.3 | 3.5 | 3.4 | 3.525 | 3.61 | 3 | 3.57 | 3.515 | 3.72 | 3.49 | 3.805 | 3.9 | 3.9849 | 3.155 | 3.725 | 3.4 | 3.3 | 3.52 | 3.55 | 3.41 | 3.105
| Std Dev. |  | 0.064083 | 0.12111 | 0 | 0.20412 | 0.040825 | 0.54772 | 0.075277 | 0.21081 | 0.057533 | 0 | 0.074944 | 0.13426 | 0.088015 | 0.065013 | 0.097707 | 0.083666 | 0.16492 | 0.075011 | 1.1148 | 0.097662 | 0.098319 | 0.14909 | 0.032711 | 0.071949 | 0.13848
| Rel.Std.Dev. |  | 1.694% | 3.303% | 0% | 5.696% | 1.243% | 15.65% | 2.225% | 5.84% | 1.591% | 0% | 2.098% | 3.832% | 2.358% | 1.861% | 2.567% | 2.173% | 4.095% | 2.356% | 26.8% | 2.875% | 2.994% | 4.204% | 0.9227% | 2.105% | 4.441%
| PDM3 |  | 6.915% | 3.618% | -1.092% | 1.263% | -7.215% | -1.092% | -4.389% | 2.017% | 2.158% | -15.22% | 0.9333% | -0.9978% | 5.502% | -1.28% | 7.574% | 8.799% | 13.8% | -10.04% | 17.56% | -4.011% | -7.215% | 0.2268% | 0.1797% | -3.4% | -11.88%
| Z-Score (Absolute) |  | 1.06 | 0.54 | 0.21 | 0.16 | 1.18 | NA | 0.73 | 0.28 | 0.31 | NA | 0.11 | 0.19 | 0.84 | 0.24 | 1.17 | 1.36 | 2.16 | 1.63 | 2.76 | 0.67 | 1.18 | 0 | 0.01 | 0.58 | 1.93
Results from laboratories 6 & 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 554 (Certified Value 0.06 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.06 | 0.0568 | 0.07 | 0.06 | 0.136 | 0.06 | 0.05 | 0.06 | 0.22 | 0.06 | 0.06 | 0.14 | 0.05 | 0.06 | 0.0745 | 0.067838 | 0.06 | 0.16 | 0.059 | 0.054 | 0.07 | 0.26 | 0.052 |  |  |  | 1
| 1 | 2 | 0.06 | 0.0561 | 0.06 | 0.06 | 0.132 | 0.06 | 0.05 | 0.06 | 0.22 | 0.06 | 0.06 | 0.16 | 0.05 | 0.06 | 0.0738 | 0.067459 | 0.06 | 0.16 | 0.059 | 0.0541 | 0.07 | 0.25 | 0.052 |  |  |  | 0
| 1 | 3 | 0.06 | 0.056 | 0.07 | 0.06 | 0.133 | 0.06 | 0.05 | 0.06 | 0.2 | 0.06 | 0.05 | 0.15 | 0.04 | 0.06 | 0.0757 | 0.05641 | 0.06 | 0.16 | 0.058 | 0.0548 | 0.07 | 0.26 | 0.052 |  |  |  | 16
| 1 | 4 | 0.06 | 0.0571 | 0.06 | 0.06 | 0.135 | 0.06 | 0.05 | 0.06 | 0.21 | 0.06 | 0.06 | 0.15 | 0.05 | 0.06 | 0.0734 | 0.068805 | 0.06 | 0.16 | 0.059 | 0.0563 | 0.07 | 0.27 | 0.052 |  |  |  | 0.059742
| 1 | 5 | 0.06 | 0.0573 | 0.07 | 0.06 | 0.136 | 0.06 | 0.05 | 0.06 | 0.23 | 0.06 | 0.06 | 0.16 | 0.05 | 0.06 | 0.075 | 0.063446 | 0.06 | 0.16 | 0.06 | 0.0541 | 0.07 | 0.27 | 0.052 |  |  |  | 112
| 1 | 6 | 0.06 | 0.0561 | 0.07 | 0.06 | 0.135 | 0.06 | 0.05 | 0.06 | 0.22 | 0.06 | 0.05 | 0.16 | 0.05 | 0.06 | 0.0738 | 0.065113 | 0.06 | 0.16 | 0.06 | 0.0575 | 0.07 | 0.28 | 0.052
| Mean |  | 0.06 | 0.056567 | 0.066667 | 0.06 | 0.1345 | 0.06 | 0.05 | 0.06 | 0.21667 | 0.06 | 0.056667 | 0.15333 | 0.048333 | 0.06 | 0.074367 | 0.064845 | 0.06 | 0.16 | 0.059167 | 0.055133 | 0.07 | 0.265 | 0.052
| Median |  | 0.06 | 0.05645 | 0.07 | 0.06 | 0.135 | 0.06 | 0.05 | 0.06 | 0.22 | 0.06 | 0.06 | 0.155 | 0.05 | 0.06 | 0.07415 | 0.066286 | 0.06 | 0.16 | 0.059 | 0.05445 | 0.07 | 0.265 | 0.052
| Std Dev. |  | 0 | 0.00057155 | 0.005164 | 0 | 0.0016432 | 0 | 7.6012e-18 | 0 | 0.010328 | 0 | 0.005164 | 0.008165 | 0.0040825 | 0 | 0.00086872 | 0.004575 | 0 | 0 | 0.00075277 | 0.0014487 | 0 | 0.010488 | 0
| Rel.Std.Dev. |  | 0% | 1.01% | 7.746% | 0% | 1.222% | 0% | 1.52e-14% | 0% | 4.767% | 0% | 9.113% | 5.325% | 8.447% | 0% | 1.168% | 7.055% | 0% | 0% | 1.272% | 2.628% | 0% | 3.958% | 0%
| PDM3 |  | 0.4314% | -5.316% | 11.59% | 0.4314% | 125.1% | 0.4314% | -16.31% | 0.4314% | 262.7% | 0.4314% | -5.148% | 156.7% | -19.1% | 0.4314% | 24.48% | 8.541% | 0.4314% | 167.8% | -0.9635% | -7.715% | 17.17% | 343.6% | -12.96%
| Z-Score (Absolute) |  | 0 | 0.48 | 0.92 | 0 | 10.32 | 0 | 1.39 | 0 | 21.71 | 0 | 0.46 | 12.93 | 1.62 | 0 | 1.99 | 0.67 | 0 | 13.86 | 0.12 | 0.67 | 1.39 | 28.4 | 1.12
| Analytical results for Nb in OREAS 554 (Certified Value 5.08 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 3.9 | 3.37 | 7 | 7 | 4 | 3 | 3.2 | 7.9 | 1.7 | < 5 | 4.5 | 6 | 6.5 | 5.2 | 4.8 | 3.72 | 5.4189 | 6.8 | 8.4114 | 2.8 | 3.57 | 5.8 | 3.2 | 7.4 |  |  | 1
| 1 | 2 | 3.7 | 4.55 | 6.5 | 6 | 4.3 | 2 | 3.3 | 7.7 | 1.7 | < 5 | 4.4 | 6.2 | 6.2 | 5.1 | 4.9 | 3.48 | 3.9851 | 6.8 | 8.518 | 2.6 | 3.81 | 5.8 | 3.6 | 7.6 |  |  | 0
| 1 | 3 | 3.2 | 4.42 | 6.5 | 6.5 | 3.5 | 1 | 3.3 | 8.3 | 1.7 | < 5 | 4.6 | 5.7 | 6.3 | 5.4 | 4.7 | 3.52 | 4.1237 | 7.2 | 8.3691 | 2.9 | 3.44 | 5.9 | 3.5 | 7.5 |  |  | 16
| 1 | 4 | 3.5 | 5.21 | 6.5 | 6.5 | 3.3 | 2 | 3.2 | 7.9 | 1.7 | < 5 | 4.5 | 6.1 | 6 | 5.1 | 4.9 | 3.59 | 3.8413 | 7.8 | 8.4574 | 2.9 | 3.79 | 6 | 3.7 | 7.6 |  |  | 5.0774
| 1 | 5 | 3.2 | 4.89 | 7 | 6.5 | 4 | 1 | 3.3 | 8 | 1.6 | < 5 | 4.9 | 5.7 | 6.3 | 5.3 | 4.9 | 3.43 | 3.7131 | 7.4 | 8.3199 | 3.1 | 3.67 | 5.8 | 3.7 | 7.6 |  |  | 113
| 1 | 6 | 3.2 | 4.82 | 6.5 | 6 | 4.7 | 1 | 3.4 | 7.5 | 1.6 | < 5 | 4.8 | 4 | 6.1 | 5.4 | 4.8 | 3.53 | 3.8698 | 7.6 | 8.4477 | 2.8 | 4.08 | 5.7 | 3.6 | 7.6
| Mean |  | 3.45 | 4.5433 | 6.6667 | 6.4167 | 3.9667 | 1.6667 | 3.2833 | 7.8833 | 1.6667 |  | 4.6167 | 5.6167 | 6.2333 | 5.25 | 4.8333 | 3.545 | 4.1587 | 7.2667 | 8.4206 | 2.85 | 3.7267 | 5.8333 | 3.55 | 7.55
| Median |  | 3.35 | 4.685 | 6.5 | 6.5 | 4 | 1.5 | 3.3 | 7.9 | 1.7 |  | 4.55 | 5.85 | 6.25 | 5.25 | 4.85 | 3.525 | 3.9275 | 7.3 | 8.4296 | 2.85 | 3.73 | 5.8 | 3.6 | 7.6
| Std Dev. |  | 0.30166 | 0.63767 | 0.2582 | 0.37639 | 0.51251 | 0.8165 | 0.075277 | 0.27142 | 0.05164 |  | 0.19408 | 0.81833 | 0.17512 | 0.13784 | 0.08165 | 0.10095 | 0.63279 | 0.41312 | 0.069933 | 0.16432 | 0.22169 | 0.10328 | 0.18708 | 0.083666
| Rel.Std.Dev. |  | 8.744% | 14.04% | 3.873% | 5.866% | 12.92% | 48.99% | 2.293% | 3.443% | 3.098% |  | 4.204% | 14.57% | 2.809% | 2.626% | 1.689% | 2.848% | 15.22% | 5.685% | 0.8305% | 5.766% | 5.949% | 1.771% | 5.27% | 1.108%
| PDM3 |  | -32.05% | -10.52% | 31.3% | 26.38% | -21.88% | -67.18% | -35.33% | 55.26% | -67.18% |  | -9.075% | 10.62% | 22.77% | 3.399% | -4.808% | -30.18% | -18.1% | 43.12% | 65.84% | -43.87% | -26.6% | 14.89% | -30.08% | 48.7%
| Z-Score (Absolute) |  | 0.65 | 0.04 | 1.14 | 1 | 0.36 | NA | 0.74 | 1.81 | 1.63 | 1.17 | 0 | 0.55 | 0.9 | 0.35 | 0.12 | 0.59 | 0.25 | 1.47 | 2.11 | 0.98 | 0.49 | 0.67 | 0.59 | 1.63
Results from laboratory 6 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 554 (Certified Value 25.3 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 19; 11: 22; 12: 24; 13: 29; 14: 30 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 11: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 14: 4A*MS |  |  | 1
2
| 1 | 1 | 01: 25; 02: 25.72; 03: 27; 04: 25.9; 05: 25.5; 06: 27.3; 09: 26.2; 10: 31.7026; 11: 26.9489; 12: 24.6; 13: 21.3; 14: 23 |  |  | 1
| 1 | 2 | 01: 24.6; 02: 25.89; 03: 26.5; 04: 25.8; 05: 24.5; 06: 28.2; 09: 25.7; 10: 32.175; 11: 28.0098; 12: 23.2; 13: 21.5; 14: 23.4 |  |  | 0
| 1 | 3 | 01: 25; 02: 26.47; 03: 26.4; 04: 26.3; 05: 24.5; 06: 29.5; 09: 26.5; 10: 32.1723; 11: 27.3244; 12: 22.8; 13: 21.7; 14: 23.2 |  |  | 16
| 1 | 4 | 01: 24.8; 02: 25.39; 03: 26.2; 04: 26.6; 05: 25.5; 06: 28.2; 09: 25.8; 10: 30.5241; 11: 26.9662; 12: 24.5; 13: 21.7; 14: 22.9 |  |  | 25.299
| 1 | 5 | 01: 24.7; 02: 25.1; 03: 26.1; 04: 26.3; 05: 25; 06: 28.3; 09: 25.9; 10: 31.1357; 11: 27.0575; 12: 22.5; 13: 22.3; 14: 22.9 |  |  | 114
| 1 | 6 | 01: 24.8; 02: 25.92; 03: 26.8; 04: 26.6; 05: 24.5; 06: 28.4; 09: 25.8; 10: 31.0506; 11: 27.7377; 12: 24.4; 13: 21.4; 14: 23.2
| Mean |  | 01: 24.8167; 02: 25.7483; 03: 26.5; 04: 26.25; 05: 24.9167; 06: 28.3167; 09: 25.9833; 10: 31.46; 11: 27.3408; 12: 23.6667; 13: 21.65; 14: 23.1
| Median |  | 01: 24.8; 02: 25.805; 03: 26.45; 04: 26.3; 05: 24.75; 06: 28.25; 09: 25.85; 10: 31.4191; 11: 27.191; 12: 23.8; 13: 21.6; 14: 23.1
| Std Dev. |  | 01: 0.16021; 02: 0.47326; 03: 0.34641; 04: 0.33912; 05: 0.4916; 06: 0.70261; 09: 0.30605; 10: 0.66727; 11: 0.44259; 12: 0.94163; 13: 0.35637; 14: 0.2
| Rel.Std.Dev. |  | 01: 0.6456%; 02: 1.838%; 03: 1.307%; 04: 1.292%; 05: 1.973%; 06: 2.481%; 09: 1.178%; 10: 2.121%; 11: 1.619%; 12: 3.979%; 13: 1.646%; 14: 0.8658%
| PDM3 |  | 01: -1.907%; 02: 1.776%; 03: 4.747%; 04: 3.759%; 05: -1.511%; 06: 11.93%; 09: 2.705%; 10: 24.35%; 11: 8.07%; 12: -6.452%; 13: -14.42%; 14: -8.692%
| Z-Score (Absolute) |  | 01: 0.56; 02: 0.06; 03: 0.34; 04: 0.21; 05: 0.51; 06: 1.31; 09: 0.06; 10: 2.99; 11: 0.79; 12: 1.18; 13: 2.25; 14: 1.48
| Analytical results for Ni in OREAS 554 (Certified Value 31.3 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 1
2
| 1 | 1 | 31 | 30.2 | 32 | 32 | 29 | 30 | 33.1 | 31.8 | 34.7 | 29 | 33.2 | 33.3 | 31.1 | 32.1 | 29.9 | 26.1 | 37.3371 | 31.6 | 33.2151 | 30 | 33.4 | 30.6 | 29.6 | 31.4 |  |  | 1
| 1 | 2 | 31 | 31.1 | 32 | 30 | 30 | 25 | 34.9 | 30 | 34.7 | 30 | 33.5 | 33.4 | 32.2 | 32.4 | 32.3 | 26.5 | 37.0252 | 30.6 | 31.9119 | 30 | 28.8 | 30.9 | 29.5 | 32.8 |  |  | 0
| 1 | 3 | 31 | 31.4 | 30 | 30 | 29 | 30 | 33.1 | 31.7 | 34 | 29 | 32.7 | 31.7 | 31.6 | 31.7 | 31 | 28.4 | 38.9359 | 32.8 | 32.4299 | 30 | 29.1 | 30.4 | 30.9 | 32.8 |  |  | 16
| 1 | 4 | 31 | 31 | 32 | 32 | 30 | 25 | 32.4 | 30.7 | 35.9 | 28 | 31.4 | 34 | 31.3 | 32.2 | 30.7 | 26 | 39.5129 | 30.7 | 31.9315 | 31 | 30.1 | 30.6 | 30.5 | 32.4 |  |  | 31.2769
| 1 | 5 | 30 | 30.6 | 32 | 30 | 30 | 25 | 34.4 | 31.3 | 34.3 | 29 | 32.1 | 32.7 | 32.8 | 32 | 30.6 | 26.5 | 39.6358 | 32.3 | 32.1154 | 31 | 28.5 | 29.9 | 30.4 | 31.5 |  |  | 115
| 1 | 6 | 31 | 30.9 | 32 | 32 | 30 | 25 | 32.5 | 30.2 | 34.2 | 28 | 32.4 | 34.6 | 31.2 | 32.7 | 29.8 | 26.7 | 38.484 | 31.3 | 32.3227 | 31 | 33.3 | 30.9 | 30.7 | 32.1
| Mean |  | 30.8333 | 30.8667 | 31.6667 | 31 | 29.6667 | 26.6667 | 33.4 | 30.95 | 34.6333 | 28.8333 | 32.55 | 33.2833 | 31.7 | 32.1833 | 30.7167 | 26.7 | 38.4885 | 31.55 | 32.3211 | 30.5 | 30.5333 | 30.55 | 30.2667 | 32.1667
| Median |  | 31 | 30.95 | 32 | 31 | 30 | 25 | 33.1 | 31 | 34.5 | 29 | 32.55 | 33.35 | 31.45 | 32.15 | 30.65 | 26.5 | 38.71 | 31.45 | 32.219 | 30.5 | 29.6 | 30.6 | 30.45 | 32.25
| Std Dev. |  | 0.40825 | 0.41793 | 0.8165 | 1.0954 | 0.5164 | 2.582 | 1.0237 | 0.76616 | 0.6802 | 0.75277 | 0.76092 | 1.0108 | 0.66933 | 0.34303 | 0.90646 | 0.87407 | 1.0984 | 0.8735 | 0.48428 | 0.54772 | 2.2474 | 0.37283 | 0.58195 | 0.61536
| Rel.Std.Dev. |  | 1.324% | 1.354% | 2.578% | 3.534% | 1.741% | 9.682% | 3.065% | 2.475% | 1.964% | 2.611% | 2.338% | 3.037% | 2.111% | 1.066% | 2.951% | 3.274% | 2.854% | 2.769% | 1.498% | 1.796% | 7.36% | 1.22% | 1.923% | 1.913%
| PDM3 |  | -1.418% | -1.312% | 1.246% | -0.8853% | -5.148% | -14.74% | 6.788% | -1.045% | 10.73% | -7.813% | 4.07% | 6.415% | 1.353% | 2.898% | -1.791% | -14.63% | 23.06% | 0.8732% | 3.339% | -2.484% | -2.377% | -2.324% | -3.23% | 2.845%
| Z-Score (Absolute) |  | 0.1 | 0.08 | 0.49 | 0.02 | 0.92 | 3.03 | 1.71 | 0.02 | 2.57 | 1.51 | 1.11 | 1.62 | 0.51 | 0.85 | 0.18 | 3.01 | 5.29 | 0.4 | 0.95 | 0.33 | 0.31 | 0.3 | 0.5 | 0.84
| Analytical results for P in OREAS 554 (Certified Value 0.031 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 3
3
| 1 | 1 | 0.03 | 0.0239 | 0.03 | 0.03 | 0.031 | 0.0315 | 0.03 | 0.038 | 0.0342 | 0.0341 | 0.032 | 0.031 | 0.028 | 0.031 | 0.032 | 0.0295 | 0.02547 | 0.0297 | 0.0316 | 0.031 | 0.0305 | 0.0308 | 0.03 | 0.0231 |  |  | 1
| 1 | 2 | 0.03 | 0.0312 | 0.03 | 0.035 | 0.03 | 0.0305 | 0.03 | 0.038 | 0.0332 | 0.0352 | 0.033 | 0.03 | 0.03 | 0.031 | 0.031 | 0.0326 | 0.02793 | 0.0303 | 0.0313 | 0.032 | 0.0313 | 0.0309 | 0.03 | 0.0236 |  |  | 0
| 1 | 3 | 0.03 | 0.0289 | 0.03 | 0.035 | 0.031 | 0.0315 | 0.03 | 0.038 | 0.0339 | 0.0347 | 0.032 | 0.03 | 0.029 | 0.029 | 0.031 | 0.0316 | 0.02811 | 0.0308 | 0.0322 | 0.032 | 0.0309 | 0.0313 | 0.03 | 0.0236 |  |  | 16
| 1 | 4 | 0.03 | 0.0277 | 0.03 | 0.035 | 0.03 | 0.031 | 0.03 | 0.038 | 0.034 | 0.0347 | 0.032 | 0.03 | 0.029 | 0.031 | 0.03 | 0.0292 | 0.02601 | 0.0304 | 0.0317 | 0.031 | 0.0319 | 0.0308 | 0.03 | 0.0079 |  |  | 0.030592
| 1 | 5 | 0.03 | 0.0289 | 0.03 | 0.035 | 0.031 | 0.0305 | 0.03 | 0.037 | 0.0351 | 0.0349 | 0.031 | 0.029 | 0.03 | 0.03 | 0.03 | 0.0318 | 0.02569 | 0.0288 | 0.0319 | 0.032 | 0.0309 | 0.0301 | 0.03 | 0.0148 |  |  | 116
| 1 | 6 | 0.03 | 0.0315 | 0.03 | 0.035 | 0.031 | 0.031 | 0.03 | 0.038 | 0.0321 | 0.0354 | 0.031 | 0.031 | 0.029 | 0.032 | 0.031 | 0.0323 | 0.0271 | 0.0304 | 0.0301 | 0.031 | 0.0322 | 0.0301 | 0.03 | 0.0236
| Mean |  | 0.03 | 0.028683 | 0.03 | 0.034167 | 0.030667 | 0.031 | 0.03 | 0.037833 | 0.03375 | 0.034833 | 0.031833 | 0.030167 | 0.029167 | 0.030667 | 0.030833 | 0.031167 | 0.026718 | 0.030067 | 0.031467 | 0.0315 | 0.031283 | 0.030667 | 0.03 | 0.019433
| Median |  | 0.03 | 0.0289 | 0.03 | 0.035 | 0.031 | 0.031 | 0.03 | 0.038 | 0.03395 | 0.0348 | 0.032 | 0.03 | 0.029 | 0.031 | 0.031 | 0.0317 | 0.026555 | 0.03035 | 0.03165 | 0.0315 | 0.0311 | 0.0308 | 0.03 | 0.02335
| Std Dev. |  | 0 | 0.0027644 | 0 | 0.0020412 | 0.0005164 | 0.00044721 | 0 | 0.00040825 | 0.0010134 | 0.00045461 | 0.00075277 | 0.00075277 | 0.00075277 | 0.0010328 | 0.00075277 | 0.0014542 | 0.0011548 | 0.00071461 | 0.00073394 | 0.00054772 | 0.00065243 | 0.0004761 | 0 | 0.0066335
| Rel.Std.Dev. |  | 0% | 9.638% | 0% | 5.974% | 1.684% | 1.443% | 0% | 1.079% | 3.003% | 1.305% | 2.365% | 2.495% | 2.581% | 3.368% | 2.441% | 4.666% | 4.322% | 2.377% | 2.332% | 1.739% | 2.086% | 1.552% | 0% | 34.13%
| PDM3 |  | -1.934% | -6.238% | -1.934% | 11.69% | 0.2452% | 1.335% | -1.934% | 23.67% | 10.32% | 13.87% | 4.059% | -1.389% | -4.658% | 0.2452% | 0.79% | 1.88% | -12.66% | -1.716% | 2.86% | 2.969% | 2.261% | 0.2452% | -1.934% | -36.48%
| Z-Score (Absolute) |  | 0.67 | 2.01 | 0.67 | 3.54 | 0 | 0.34 | 0.67 | 7.25 | 3.12 | 4.21 | 1.18 | 0.51 | 1.52 | 0 | 0.17 | 0.51 | 3.99 | 0.61 | 0.81 | 0.84 | 0.62 | 0 | 0.67 | 11.38
| Analytical results for Pb in OREAS 554 (Certified Value 16.6 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 16.2 | 18.7 | 16 | 16 | 17 | 18 | 16.9 | 17.6 | 17.9 | 10 | 16.7 | 17 | 15.2 | 15.4 | 18 | 19 | 17 | 15.1 | 15.686 | 17 | 16.66 | 17 | 17 | 16.9 |  |  | 1
| 1 | 2 | 16.1 | 18.1 | 17 | 16 | 16 | 17 | 16.5 | 17.6 | 17.4 | 10 | 17 | 16.8 | 15.8 | 15.6 | 18.9 | 12 | 16.6 | 18 | 15.9092 | 16 | 16.56 | 17.1 | 16 | 16.9 |  |  | 0
| 1 | 3 | 16.2 | 18.1 | 16 | 17 | 17 | 17 | 17.1 | 18.6 | 18 | 10 | 16.9 | 16 | 14.5 | 15.4 | 17.7 | 13 | 16.7 | 15.7 | 16.7212 | 17 | 17.05 | 17.4 | 17 | 17.5 |  |  | 16
| 1 | 4 | 16.3 | 18.1 | 16 | 15 | 17 | 17 | 15.8 | 17.3 | 17.3 | 11 | 16.5 | 16.7 | 15.1 | 15.3 | 17.8 | 11 | 17.5 | 15.3 | 15.5333 | 17 | 16.86 | 17.3 | 15 | 17.2 |  |  | 16.5952
| 1 | 5 | 16 | 18.2 | 17 | 15 | 17 | 17 | 16.4 | 17.6 | 18.1 | 12 | 17.4 | 16.5 | 14.2 | 15.1 | 17.4 | 12 | 17.9 | 15.7 | 16.1254 | 17 | 16.86 | 17.2 | 16 | 16.4 |  |  | 117
| 1 | 6 | 15.9 | 18 | 16 | 15 | 16 | 17 | 16.3 | 17.3 | 16.7 | 12 | 16.7 | 18.6 | 14 | 15.5 | 17.4 | 11 | 16.8 | 14.9 | 15.9552 | 17 | 16.46 | 17.1 | 15 | 16.6
| Mean |  | 16.1167 | 18.2 | 16.3333 | 15.6667 | 16.6667 | 17.1667 | 16.5 | 17.6667 | 17.5667 | 10.8333 | 16.8667 | 16.9333 | 14.8 | 15.3833 | 17.8667 | 13 | 17.0833 | 15.7833 | 15.9884 | 16.8333 | 16.7417 | 17.1833 | 16 | 16.9167
| Median |  | 16.15 | 18.1 | 16 | 15.5 | 17 | 17 | 16.45 | 17.6 | 17.65 | 10.5 | 16.8 | 16.75 | 14.8 | 15.4 | 17.75 | 12 | 16.9 | 15.5 | 15.9322 | 17 | 16.76 | 17.15 | 16 | 16.9
| Std Dev. |  | 0.1472 | 0.25298 | 0.5164 | 0.8165 | 0.5164 | 0.40825 | 0.46043 | 0.48028 | 0.53541 | 0.98319 | 0.31411 | 0.88468 | 0.68411 | 0.17224 | 0.55737 | 3.0332 | 0.51153 | 1.1321 | 0.41514 | 0.40825 | 0.22004 | 0.1472 | 0.89443 | 0.39707
| Rel.Std.Dev. |  | 0.9133% | 1.39% | 3.162% | 5.212% | 3.098% | 2.378% | 2.791% | 2.719% | 3.048% | 9.076% | 1.862% | 5.225% | 4.622% | 1.12% | 3.12% | 23.33% | 2.994% | 7.173% | 2.597% | 2.425% | 1.314% | 0.8566% | 5.59% | 2.347%
| PDM3 |  | -2.883% | 9.671% | -1.578% | -5.595% | 0.4309% | 3.444% | -0.5734% | 6.457% | 5.854% | -34.72% | 1.636% | 2.038% | -10.82% | -7.302% | 7.662% | -21.66% | 2.942% | -4.892% | -3.656% | 1.435% | 0.8829% | 3.544% | -3.586% | 1.937%
| Z-Score (Absolute) |  | 0.61 | 1.56 | 0.39 | 1.08 | 0.04 | 0.48 | 0.21 | 1 | 0.9 | 6.13 | 0.17 | 0.24 | 1.99 | 1.38 | 1.21 | 3.87 | 0.4 | 0.96 | 0.75 | 0.13 | 0.04 | 0.5 | 0.74 | 0.22
| Analytical results for Pr in OREAS 554 (Certified Value 6.86 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 22; 11: 24; 12: 29; 13: 30 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*MS |  |  | 2
3
| 1 | 1 | 01: 7.06; 02: 7.06; 03: 7.22; 04: 7.1; 05: 6.6; 06: 7.4; 09: 6.66; 10: 7.0787; 11: 6.5; 12: 5.47; 13: 6.2 |  |  | 1
| 1 | 2 | 01: 7; 02: 7.17; 03: 7.04; 04: 6.9; 05: 6.3; 06: 7.3; 09: 6.54; 10: 7.2418; 11: 6.1; 12: 5.52; 13: 6.5 |  |  | 0
| 1 | 3 | 01: 6.99; 02: 7.23; 03: 7.28; 04: 7; 05: 6.4; 06: 7.9; 09: 6.38; 10: 7.4002; 11: 6.1; 12: 5.57; 13: 6.3 |  |  | 16
| 1 | 4 | 01: 6.9; 02: 7.26; 03: 7.04; 04: 7.15; 05: 6.8; 06: 7.3; 09: 6.71; 10: 7.4017; 11: 6.6; 12: 5.61; 13: 6.5 |  |  | 6.8615
| 1 | 5 | 01: 7.04; 02: 7.12; 03: 6.96; 04: 7.15; 05: 6.6; 06: 7.5; 09: 6.57; 10: 7.2428; 11: 5.9; 12: 5.71; 13: 6.2 |  |  | 118
| 1 | 6 | 01: 7.04; 02: 7.05; 03: 7.28; 04: 7.1; 05: 6.4; 06: 7.3; 09: 6.46; 10: 7.2027; 11: 6.1; 12: 5.48; 13: 6.4
| Mean |  | 01: 7.005; 02: 7.1483; 03: 7.1367; 04: 7.0667; 05: 6.5167; 06: 7.45; 09: 6.5533; 10: 7.2613; 11: 6.2167; 12: 5.56; 13: 6.35
| Median |  | 01: 7.02; 02: 7.145; 03: 7.13; 04: 7.1; 05: 6.5; 06: 7.35; 09: 6.555; 10: 7.2423; 11: 6.1; 12: 5.545; 13: 6.35
| Std Dev. |  | 01: 0.057879; 02: 0.087044; 03: 0.13995; 04: 0.098319; 05: 0.18348; 06: 0.23452; 09: 0.12258; 10: 0.12373; 11: 0.27142; 12: 0.090774; 13: 0.13784
| Rel.Std.Dev. |  | 01: 0.8263%; 02: 1.218%; 03: 1.961%; 04: 1.391%; 05: 2.816%; 06: 3.148%; 09: 1.871%; 10: 1.704%; 11: 4.366%; 12: 1.633%; 13: 2.171%
| PDM3 |  | 01: 2.092%; 02: 4.181%; 03: 4.011%; 04: 2.991%; 05: -5.025%; 06: 8.577%; 09: -4.491%; 10: 5.827%; 11: -9.397%; 12: -18.97%; 13: -7.454%
| Z-Score (Absolute) |  | 01: 0; 02: 0.22; 03: 0.2; 04: 0.09; 05: 0.74; 06: 0.67; 09: 0.68; 10: 0.39; 11: 1.19; 12: 2.19; 13: 0.99
| Analytical results for Rb in OREAS 554 (Certified Value 88 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
1
| 1 | 1 | 92 | 88.66 | 91 | 89.4 | 85.2 | 86.5 | 84.8 | 88.3 | 89.2 | 91.4 | 91.6 | 90.8 | 90.4 | 84.2 | 90 | 81.2488 | 88.5 | 94.3 | 82.2464 | 78.9 | 90.9 | 90.5 | 87.3 | 80.6 |  |  | 1
| 1 | 2 | 89.7 | 88.36 | 90.2 | 87 | 84.3 | 85 | 85.4 | 87 | 89.2 | 92.2 | 92.3 | 92 | 92.8 | 88 | 92 | 82.5371 | 86.5 | 96.5 | 84.7063 | 78.5 | 85.6 | 89.8 | 87.6 | 83.8 |  |  | 0
| 1 | 3 | 89.9 | 88.05 | 91.6 | 88.2 | 84.9 | 85 | 88.2 | 92.2 | 89.2 | 90.6 | 89.2 | 91.1 | 89.7 | 85.4 | 92 | 81.4382 | 87.6 | 99.1 | 81.8282 | 79.7 | 85.2 | 91 | 88.4 | 82.7 |  |  | 16
| 1 | 4 | 91.1 | 88.06 | 89.4 | 89.6 | 85.9 | 88 | 85.1 | 81.4 | 90.3 | 89.8 | 89.7 | 91.5 | 92.8 | 86.9 | 91 | 80.2086 | 86.2 | 100.7 | 83.1923 | 80.6 | 85.7 | 90.8 | 87.8 | 83 |  |  | 87.9223
| 1 | 5 | 88.8 | 88.9 | 91.8 | 88 | 85.9 | 87.5 | 85.7 | 87.2 | 87.7 | 87.8 | 91.5 | 92.8 | 89.9 | 88.2 | 94 | 79.8138 | 86.7 | 95.7 | 83.3969 | 82 | 89.1 | 91.3 | 90.9 | 82 |  |  | 119
| 1 | 6 | 89.3 | 89.84 | 92.4 | 88.8 | 85 | 83.5 | 85.2 | 89.6 | 87.9 | 88.6 | 92.3 | 93.1 | 92.2 | 82.3 | 93 | 80.7294 | 88.9 | 94.2 | 83.2023 | 80.9 | 90.8 | 91.4 | 87.4 | 84.6
| Mean |  | 90.1333 | 88.645 | 91.0667 | 88.5 | 85.2 | 85.9167 | 85.7333 | 87.6167 | 88.9167 | 90.0667 | 91.1 | 91.8833 | 91.3 | 85.8333 | 92 | 80.996 | 87.4 | 96.75 | 83.0954 | 80.1 | 87.8833 | 90.8 | 88.2333 | 82.7833
| Median |  | 89.8 | 88.51 | 91.3 | 88.5 | 85.1 | 85.75 | 85.3 | 87.75 | 89.2 | 90.2 | 91.55 | 91.75 | 91.3 | 86.15 | 92 | 80.9891 | 87.15 | 96.1 | 83.1973 | 80.15 | 87.4 | 90.9 | 87.7 | 82.85
| Std Dev. |  | 1.1944 | 0.67391 | 1.1075 | 0.96954 | 0.61968 | 1.7151 | 1.2453 | 3.5947 | 0.96626 | 1.6717 | 1.3312 | 0.92394 | 1.4588 | 2.314 | 1.4142 | 0.9719 | 1.1171 | 2.638 | 1.0033 | 1.3161 | 2.6933 | 0.58992 | 1.3633 | 1.3977
| Rel.Std.Dev. |  | 1.325% | 0.7602% | 1.216% | 1.096% | 0.7273% | 1.996% | 1.452% | 4.103% | 1.087% | 1.856% | 1.461% | 1.006% | 1.598% | 2.696% | 1.537% | 1.2% | 1.278% | 2.727% | 1.207% | 1.643% | 3.065% | 0.6497% | 1.545% | 1.688%
| PDM3 |  | 2.515% | 0.8219% | 3.576% | 0.657% | -3.096% | -2.281% | -2.49% | -0.3477% | 1.131% | 2.439% | 3.614% | 4.505% | 3.842% | -2.376% | 4.638% | -7.878% | -0.5941% | 10.04% | -5.49% | -8.897% | -0.04438% | 3.273% | 0.3537% | -5.845%
| Z-Score (Absolute) |  | 0.46 | 0.07 | 0.7 | 0.03 | 0.83 | 0.64 | 0.69 | 0.2 | 0.14 | 0.44 | 0.71 | 0.92 | 0.77 | 0.66 | 0.95 | 1.92 | 0.25 | 2.19 | 1.38 | 2.16 | 0.13 | 0.64 | 0.03 | 1.46
| Analytical results for Re in OREAS 554 (Certified Value 0.044 ppm) | [27 empty] | Certified
Re | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 05; 04: 06; 05: 10; 06: 13; 09: 14; 10: 15; 11: 16; 13: 17; 14: 23; 15: 24; 16: 28 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES; 04: 4A*OES/MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*OES/MS; 10: 4A*OES/MS; 11: 4A*OES/MS; 13: 4A*OES/MS; 14: 4A*MS; 15: 4A*MS; 16: 4A*OES/MS |  |  | 3
3
| 1 | 1 | 01: 0.038; 02: < 0.1; 03: < 0.05; 04: 0.05; 05: 0.043; 06: 0.046; 09: 0.047; 10: 0.044; 11: 0.046; 13: 0.057; 14: < 0.05; 15: 0.038; 16: 0.045 |  |  | 1
| 1 | 2 | 01: 0.054; 02: < 0.1; 03: < 0.05; 04: 0.05; 05: 0.042; 06: 0.046; 09: 0.051; 10: 0.04; 11: 0.044; 13: 0.046; 14: < 0.05; 15: 0.041; 16: 0.043 |  |  | 0
| 1 | 3 | 01: 0.041; 02: < 0.1; 03: < 0.05; 04: 0.05; 05: 0.047; 06: 0.048; 09: 0.043; 10: 0.045; 11: 0.044; 13: 0.04; 14: < 0.05; 15: 0.039; 16: 0.043 |  |  | 16
| 1 | 4 | 01: 0.045; 02: < 0.1; 03: < 0.05; 04: 0.05; 05: 0.044; 06: 0.05; 09: 0.041; 10: 0.045; 11: 0.044; 13: 0.039; 14: < 0.05; 15: 0.036; 16: 0.043 |  |  | 0.043952
| 1 | 5 | 01: 0.045; 02: < 0.1; 03: < 0.05; 04: 0.05; 05: 0.045; 06: 0.044; 09: 0.039; 10: 0.047; 11: 0.044; 13: 0.051; 14: < 0.05; 15: 0.035; 16: 0.044 |  |  | 120
| 1 | 6 | 01: 0.041; 02: < 0.1; 03: < 0.05; 04: 0.05; 05: 0.044; 06: 0.042; 09: 0.042; 10: 0.043; 11: 0.043; 13: 0.03; 14: < 0.05; 15: 0.035; 16: 0.043
| Mean |  | 01: 0.044; 04: 0.05; 05: 0.044167; 06: 0.046; 09: 0.043833; 10: 0.044; 11: 0.044167; 13: 0.043833; 15: 0.037333; 16: 0.0435
| Median |  | 01: 0.043; 04: 0.05; 05: 0.044; 06: 0.046; 09: 0.0425; 10: 0.0445; 11: 0.044; 13: 0.043; 15: 0.037; 16: 0.043
| Std Dev. |  | 01: 0.0055857; 04: 7.6012e-18; 05: 0.0017224; 06: 0.0028284; 09: 0.0044008; 10: 0.0023664; 11: 0.00098319; 13: 0.0095795; 15: 0.0024221; 16: 0.00083666
| Rel.Std.Dev. |  | 01: 12.69%; 04: 1.52e-14%; 05: 3.9%; 06: 6.149%; 09: 10.04%; 10: 5.378%; 11: 2.226%; 13: 21.85%; 15: 6.488%; 16: 1.923%
| PDM3 |  | 01: 0.1083%; 04: 13.76%; 05: 0.4875%; 06: 4.659%; 09: -0.2709%; 10: 0.1083%; 11: 0.4875%; 13: -0.2709%; 15: -15.06%; 16: -1.029%
| Z-Score (Absolute) |  | 01: 0.17; 02: 12.31; 03: 38.27; 04: NA; 05: 0.51; 06: 4.21; 09: 0.17; 10: 0.17; 11: 0.51; 13: 0.17; 14: 38.27; 15: 13.32; 16: 0.84
| Analytical results for S in OREAS 554 (Certified Value 0.778 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.84 | 0.77 | 0.82 | 0.805 | 0.842 | 0.765 | 0.69 | 0.79 | 0.75 | 0.78 | 0.78 | 0.78 | 0.77 | 0.76 | 0.78 | NA | 0.71291 | 0.76 | 0.754 | 0.7479 | 0.79 | 0.79 | 0.842 |  |  |  | 1
| 1 | 2 | 0.84 | 0.81 | 0.795 | 0.805 | 0.823 | 0.77 | 0.7 | 0.78 | 0.77 | 0.78 | 0.81 | 0.76 | 0.81 | 0.75 | 0.77 | NA | 0.72801 | 0.76 | 0.757 | 0.7551 | 0.79 | 0.78 | 0.856 |  |  |  | 0
| 1 | 3 | 0.84 | 0.76 | 0.79 | 0.8 | 0.771 | 0.765 | 0.7 | 0.78 | 0.78 | 0.77 | 0.8 | 0.76 | 0.8 | 0.74 | 0.77 | NA | 0.6827 | 0.81 | 0.766 | 0.7533 | 0.79 | 0.79 | 0.862 |  |  |  | 16
| 1 | 4 | 0.84 | 0.8 | 0.805 | 0.795 | 0.768 | 0.76 | 0.68 | 0.78 | 0.79 | 0.77 | 0.78 | 0.74 | 0.78 | 0.74 | 0.73 | NA | 0.69353 | 0.79 | 0.755 | 0.7722 | 0.8 | 0.77 | 0.862 |  |  |  | 0.77837
| 1 | 5 | 0.85 | 0.78 | 0.81 | 0.775 | 0.844 | 0.745 | 0.68 | 0.77 | 0.78 | 0.77 | 0.77 | 0.76 | 0.78 | 0.73 | 0.77 | NA | 0.69248 | 0.82 | 0.756 | 0.7614 | 0.78 | 0.81 | 0.856 |  |  |  | 121
| 1 | 6 | 0.87 | 0.78 | 0.815 | 0.785 | 0.828 | 0.755 | 0.68 | 0.77 | 0.78 | 0.78 | 0.76 | 0.83 | 0.79 | 0.77 | 0.75 | NA | 0.69476 | 0.76 | 0.753 | 0.7686 | 0.76 | 0.82 | 0.865
| Mean |  | 0.84667 | 0.78333 | 0.80583 | 0.79417 | 0.81267 | 0.76 | 0.68833 | 0.77833 | 0.775 | 0.775 | 0.78333 | 0.77167 | 0.78833 | 0.74833 | 0.76167 |  | 0.70073 | 0.78333 | 0.75683 | 0.75975 | 0.785 | 0.79333 | 0.85717
| Median |  | 0.84 | 0.78 | 0.8075 | 0.7975 | 0.8255 | 0.7625 | 0.685 | 0.78 | 0.78 | 0.775 | 0.78 | 0.76 | 0.785 | 0.745 | 0.77 |  | 0.69414 | 0.775 | 0.7555 | 0.75825 | 0.79 | 0.79 | 0.859
| Std Dev. |  | 0.012111 | 0.018619 | 0.011583 | 0.012007 | 0.034396 | 0.0089443 | 0.0098319 | 0.0075277 | 0.013784 | 0.0054772 | 0.018619 | 0.031252 | 0.01472 | 0.01472 | 0.018348 |  | 0.016572 | 0.027325 | 0.0047081 | 0.0093776 | 0.013784 | 0.018619 | 0.0082563
| Rel.Std.Dev. |  | 1.43% | 2.377% | 1.437% | 1.512% | 4.232% | 1.177% | 1.428% | 0.9672% | 1.779% | 0.7067% | 2.377% | 4.05% | 1.867% | 1.967% | 2.409% |  | 2.365% | 3.488% | 0.6221% | 1.234% | 1.756% | 2.347% | 0.9632%
| PDM3 |  | 8.775% | 0.6382% | 3.529% | 2.03% | 4.407% | -2.36% | -11.57% | -0.004164% | -0.4324% | -0.4324% | 0.6382% | -0.8607% | 1.281% | -3.858% | -2.145% |  | -9.974% | 0.6382% | -2.766% | -2.392% | 0.8523% | 1.923% | 10.12%
| Z-Score (Absolute) |  | 2.73 | 0.1 | 1.04 | 0.55 | 1.32 | 0.86 | 3.84 | 0.1 | 0.24 | 0.24 | 0.1 | 0.38 | 0.31 | 1.35 | 0.8 | NA | 3.32 | 0.1 | 1 | 0.87 | 0.17 | 0.52 | 3.17
| Analytical results for Sb in OREAS 554 (Certified Value 2.22 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.26 | 2.35 | 2.2 | 2.3 | 1.9 | 1 | 2 | 2.5 | 2.58 | < 5 | 2.3 | 2.38 | 2.21 | 2.19 | 2.25 | 2.61 | 2.3804 | 1.98 | 2.28 | 2.1789 | 1.9 | 1.95 | 2.3 | 2.06 |  |  | 1
| 1 | 2 | 2.22 | 2.31 | 2.2 | 2.2 | 1.8 | 1 | 2.1 | 2.4 | 2.53 | < 5 | 2.33 | 2.42 | 2.28 | 2.27 | 2.28 | 2.58 | 2.5131 | 2.08 | 2.28 | 2.255 | 1.8 | 2.02 | 2.2 | 2.11 |  |  | 0
| 1 | 3 | 2.18 | 2.32 | 2 | 2.3 | 1.9 | 1 | 2 | 2.6 | 2.52 | < 5 | 2.3 | 2.33 | 2.19 | 2.18 | 2.19 | 2.53 | 2.437 | 1.98 | 2.29 | 2.2136 | 1.8 | 1.98 | 2.3 | 2.03 |  |  | 16
| 1 | 4 | 2.26 | 2.36 | 2 | 2.2 | 1.8 | 1 | 2 | 2.5 | 2.37 | < 5 | 2.21 | 2.44 | 2.23 | 2.25 | 2.32 | 2.47 | 2.4312 | 1.96 | 2.27 | 2.2386 | 1.8 | 1.91 | 2.3 | 2.06 |  |  | 2.2204
| 1 | 5 | 2.2 | 2.28 | 2.2 | 2.3 | 2 | 1 | 1.7 | 2.5 | 2.48 | < 5 | 2.28 | 2.4 | 2.23 | 2.21 | 2.27 | 2.55 | 2.2719 | 2.02 | 2.38 | 2.232 | 1.8 | 1.91 | 2.2 | 2.17 |  |  | 122
| 1 | 6 | 2.32 | 2.44 | 2.2 | 2.2 | 1.9 | 1 | 2 | 2.4 | 2.46 | < 5 | 2.26 | 2.39 | 2.23 | 2.24 | 2.2 | 2.57 | 2.3773 | 2.04 | 2.31 | 2.2223 | 1.8 | 2.13 | 2.3 | 2.11
| Mean |  | 2.24 | 2.3433 | 2.1333 | 2.25 | 1.8833 | 1 | 1.9667 | 2.4833 | 2.49 |  | 2.28 | 2.3933 | 2.2283 | 2.2233 | 2.2517 | 2.5517 | 2.4018 | 2.01 | 2.3017 | 2.2234 | 1.8167 | 1.9833 | 2.2667 | 2.09
| Median |  | 2.24 | 2.335 | 2.2 | 2.25 | 1.9 | 1 | 2 | 2.5 | 2.5 |  | 2.29 | 2.395 | 2.23 | 2.225 | 2.26 | 2.56 | 2.4058 | 2 | 2.285 | 2.2271 | 1.8 | 1.965 | 2.3 | 2.085
| Std Dev. |  | 0.050596 | 0.055377 | 0.10328 | 0.054772 | 0.075277 | 0 | 0.13663 | 0.075277 | 0.072111 |  | 0.041473 | 0.037771 | 0.029944 | 0.03559 | 0.049565 | 0.048339 | 0.08051 | 0.045166 | 0.040702 | 0.026004 | 0.040825 | 0.083347 | 0.05164 | 0.0502
| Rel.Std.Dev. |  | 2.259% | 2.363% | 4.841% | 2.434% | 3.997% | 0% | 6.947% | 3.031% | 2.896% |  | 1.819% | 1.578% | 1.344% | 1.601% | 2.201% | 1.894% | 3.352% | 2.247% | 1.768% | 1.17% | 2.247% | 4.202% | 2.278% | 2.402%
| PDM3 |  | 0.8812% | 5.535% | -3.923% | 1.332% | -15.18% | -54.96% | -11.43% | 11.84% | 12.14% |  | 2.683% | 7.787% | 0.3558% | 0.1306% | 1.407% | 14.92% | 8.169% | -9.477% | 3.658% | 0.1333% | -18.18% | -10.68% | 2.082% | -5.874%
| Z-Score (Absolute) |  | 0.05 | 0.44 | 0.55 | 0 | 1.72 | NA | 1.33 | 1.1 | 1.13 | 1.18 | 0.14 | 0.67 | 0.1 | 0.13 | 0.01 | 1.42 | 0.71 | 1.13 | 0.24 | 0.13 | 2.04 | 1.25 | 0.08 | 0.75
Results from laboratory 6 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 554 (Certified Value 8.1 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 8 | 8.1 | 8 | 9 | 8 | 8 | 9 | 8 | 8.5 | 7 | 8.3 | 8.1 | 8.4 | 8 | 7.6 | 8.13 | 8.134 | 8.9 | 8.4 | 8.1857 | 8 | 7.9 | 7.9 | 7.3 | 7.2 |  | 1
| 1 | 2 | 7.9 | 8.1 | 8 | 8 | 8 | 8 | 8 | 8 | 8.3 | 7 | 8.2 | 8.2 | 8.7 | 8.2 | 7.9 | 8.14 | 8.3971 | 8.9 | 9 | 8.422 | 8 | 7.9 | 7.9 | 8.2 | 7.5 |  | 0
| 1 | 3 | 8 | 7.9 | 8 | 9 | 8 | 8 | 9 | 8 | 8.5 | 6 | 8.2 | 8 | 8.4 | 7.9 | 7.7 | 8.29 | 8.4376 | 8.9 | 9.3 | 8.2186 | 8 | 7.9 | 7.9 | 8.3 | 7.4 |  | 16
| 1 | 4 | 8 | 7.8 | 8 | 9 | 8 | 8 | 8 | 8 | 8.5 | 6 | 8 | 8 | 8.4 | 8.2 | 7.7 | 8.04 | 8.4743 | 9.1 | 8.9 | 8.2639 | 8 | 8.2 | 7.9 | 8.2 | 7.3 |  | 8.1008
| 1 | 5 | 7.9 | 7.9 | 8 | 9 | 8 | 8 | 8 | 8 | 8.6 | 6 | 8.1 | 8.1 | 8.4 | 8 | 7.7 | 8.29 | 7.8573 | 8.8 | 9.1 | 8.1811 | 8 | 7.9 | 8 | 7.8 | 7.3 |  | 123
| 1 | 6 | 7.9 | 8 | 8 | 8 | 8 | 8 | 8 | 8 | 8.4 | 6 | 8.2 | 8.9 | 8.6 | 8.1 | 7.4 | 8.11 | 7.9389 | 9.1 | 8.9 | 8.1773 | 8 | 8.2 | 7.8 | 8.2 | 7.5
| Mean |  | 7.95 | 7.9667 | 8 | 8.6667 | 8 | 8 | 8.3333 | 8 | 8.4667 | 6.3333 | 8.1667 | 8.2167 | 8.4833 | 8.0667 | 7.6667 | 8.1667 | 8.2065 | 8.95 | 8.9333 | 8.2414 | 8 | 8 | 7.9 | 8 | 7.3667
| Median |  | 7.95 | 7.95 | 8 | 9 | 8 | 8 | 8 | 8 | 8.5 | 6 | 8.2 | 8.1 | 8.4 | 8.05 | 7.7 | 8.135 | 8.2655 | 8.9 | 8.95 | 8.2021 | 8 | 7.9 | 7.9 | 8.2 | 7.35
| Std Dev. |  | 0.054772 | 0.12111 | 0 | 0.5164 | 0 | 0 | 0.5164 | 0 | 0.10328 | 0.5164 | 0.10328 | 0.34303 | 0.13292 | 0.12111 | 0.1633 | 0.10172 | 0.26843 | 0.12247 | 0.30111 | 0.094311 | 0 | 0.15492 | 0.063246 | 0.38471 | 0.12111
| Rel.Std.Dev. |  | 0.689% | 1.52% | 0% | 5.958% | 0% | 0% | 6.197% | 0% | 1.22% | 8.154% | 1.265% | 4.175% | 1.567% | 1.501% | 2.13% | 1.246% | 3.271% | 1.368% | 3.371% | 1.144% | 0% | 1.936% | 0.8006% | 4.809% | 1.644%
| PDM3 |  | -1.861% | -1.655% | -1.244% | 6.986% | -1.244% | -1.244% | 2.871% | -1.244% | 4.517% | -21.82% | 0.8137% | 1.431% | 4.723% | -0.4208% | -5.359% | 0.8137% | 1.306% | 10.48% | 10.28% | 1.736% | -1.244% | -1.244% | -2.478% | -1.244% | -9.062%
| Z-Score (Absolute) |  | 0.82 | 0.74 | 0.57 | NA | NA | NA | NA | NA | 1.72 | NA | 0.25 | 0.49 | 1.8 | 0.25 | 2.21 | 0.25 | 0.44 | 4.09 | 4.01 | 0.61 | NA | 0.57 | 1.06 | 0.57 | 3.68
Results from laboratories 4, 5, 6, 9, 10 & 23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Se in OREAS 554 (Certified Value 5.29 ppm) | [27 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 5 | 6.4 | 5 | < 5 | 5 | 5 | < 5 | 5.3 | 7.5 | 6 | 6 | 5 | 7 | 6 | 5 | 4.9 | 4.5027 | 4.5 | 4 | 6 | 5.6 | 6 | 5 |  |  |  | 1
| 1 | 2 | 5 | 6.5 | < 5 | < 5 | 5 | < 5 | 6 | 5 | 7.5 | 5 | 6 | 5 | 7 | 6 | 6 | 5.2 | 4.7804 | 5.2 | 4 | 5 | 5 | 5 | 4 |  |  |  | 0
| 1 | 3 | 4 | 6.6 | < 5 | < 5 | 5 | < 5 | 7 | 6 | 7.8 | 5 | 6 | 6 | 5 | 5 | 4 | 4.9 | 4.3549 | 5.4 | 5 | 5 | 5.5 | 5 | 4 |  |  |  | 16
| 1 | 4 | 5 | 6.2 | 5 | < 5 | 5 | 10 | 5 | 6.1 | 7.6 | 5 | 5 | 5 | 7 | 5 | 6 | 5 | 4.2198 | 5.3 | 5 | 6 | 5.2 | 5 | 6 |  |  |  | 5.2926
| 1 | 5 | 5 | 5.7 | < 5 | < 5 | 5 | 10 | 6 | 4.4 | 7.1 | 5 | 7 | 5 | 6 | 5 | 6 | 5 | 4.3059 | 4.8 | 4 | 6 | 5 | 5 | 6 |  |  |  | 124
| 1 | 6 | 5 | 5.9 | 5 | < 5 | 5 | < 5 | 6 | 4.4 | 7.2 | 5 | 5 | 6 | 7 | 5 | 7 | 4.7 | 4.5443 | 5.3 | 4 | 6 | 5.4 | 5 | 4
| Mean |  | 4.8333 | 6.2167 | 5 |  | 5 | 8.3333 | 6 | 5.2 | 7.45 | 5.1667 | 5.8333 | 5.3333 | 6.5 | 5.3333 | 5.6667 | 4.95 | 4.4513 | 5.0833 | 4.3333 | 5.6667 | 5.2833 | 5.1667 | 4.8333
| Median |  | 5 | 6.3 | 5 |  | 5 | 10 | 6 | 5.15 | 7.5 | 5 | 6 | 5 | 7 | 5 | 6 | 4.95 | 4.4288 | 5.25 | 4 | 6 | 5.3 | 5 | 4.5
| Std Dev. |  | 0.40825 | 0.35449 | 0 |  | 0 | 2.8868 | 0.70711 | 0.74565 | 0.25884 | 0.40825 | 0.75277 | 0.5164 | 0.83666 | 0.5164 | 1.0328 | 0.16432 | 0.20181 | 0.35449 | 0.5164 | 0.5164 | 0.25626 | 0.40825 | 0.98319
| Rel.Std.Dev. |  | 8.447% | 5.702% | 0% |  | 0% | 34.64% | 11.79% | 14.34% | 3.474% | 7.902% | 12.9% | 9.682% | 12.87% | 9.682% | 18.23% | 3.32% | 4.534% | 6.974% | 11.92% | 9.113% | 4.85% | 7.902% | 20.34%
| PDM3 |  | -8.677% | 17.46% | -5.528% |  | -5.528% | 57.45% | 13.37% | -1.749% | 40.76% | -2.379% | 10.22% | 0.7703% | 22.81% | 0.7703% | 7.068% | -6.473% | -15.89% | -3.953% | -18.12% | 7.068% | -0.1745% | -2.379% | -8.677%
| Z-Score (Absolute) |  | 0.67 | 1.87 | 2.67 | 4.97 | 0.37 | 0.4 | 0.4 | 0 | 4.14 | 0.06 | 1.16 | 0.25 | 2.39 | 0.25 | 0.86 | 0.46 | 1.38 | 0.21 | 1.59 | 0.86 | 0.15 | 0.06 | 0.67
| Analytical results for Sm in OREAS 554 (Certified Value 4.82 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 22; 11: 24; 12: 29; 13: 30 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*MS |  |  |  | 2
3
| 1 | 1 | 01: 4.6; 02: 4.96; 03: 4.95; 04: 5.05; 05: 5; 06: 4.7; 09: 5.3; 10: 4.5191; 11: 4.4; 12: 4.1; 13: 4.2 |  |  |  | 1
| 1 | 2 | 01: 4.6; 02: 4.83; 03: 4.8; 04: 4.8; 05: 4.5; 06: 4.6; 09: 5.1; 10: 4.4434; 11: 4.1; 12: 4.2; 13: 4.3 |  |  |  | 0
| 1 | 3 | 01: 4.7; 02: 5; 03: 4.8; 04: 4.95; 05: 4.7; 06: 5.6; 09: 5.1; 10: 4.8118; 11: 4.2; 12: 4.2; 13: 4.2 |  |  |  | 16
| 1 | 4 | 01: 4.5; 02: 4.85; 03: 4.8; 04: 4.9; 05: 5.1; 06: 4.9; 09: 5; 10: 4.7197; 11: 4.3; 12: 4.2; 13: 4.2 |  |  |  | 4.8186
| 1 | 5 | 01: 4.5; 02: 4.91; 03: 4.8; 04: 4.9; 05: 5.1; 06: 4.9; 09: 5; 10: 4.6683; 11: 4.4; 12: 4.2; 13: 4.2 |  |  |  | 125
| 1 | 6 | 01: 4.7; 02: 4.93; 03: 4.8; 04: 4.85; 05: 4.8; 06: 4.4; 09: 5; 10: 4.8985; 11: 4.2; 12: 4.1; 13: 4.2
| Mean |  | 01: 4.6; 02: 4.9133; 03: 4.825; 04: 4.9083; 05: 4.8667; 06: 4.85; 09: 5.0833; 10: 4.6768; 11: 4.2667; 12: 4.1667; 13: 4.2167
| Median |  | 01: 4.6; 02: 4.92; 03: 4.8; 04: 4.9; 05: 4.9; 06: 4.8; 09: 5.05; 10: 4.694; 11: 4.25; 12: 4.2; 13: 4.2
| Std Dev. |  | 01: 0.089443; 02: 0.064704; 03: 0.061237; 04: 0.08612; 05: 0.24221; 06: 0.41352; 09: 0.1169; 10: 0.1724; 11: 0.12111; 12: 0.05164; 13: 0.040825
| Rel.Std.Dev. |  | 01: 1.944%; 02: 1.317%; 03: 1.269%; 04: 1.755%; 05: 4.977%; 06: 8.526%; 09: 2.3%; 10: 3.686%; 11: 2.838%; 12: 1.239%; 13: 0.9682%
| PDM3 |  | 01: -4.536%; 02: 1.967%; 03: 0.1337%; 04: 1.863%; 05: 0.9984%; 06: 0.6525%; 09: 5.495%; 10: -2.942%; 11: -11.45%; 12: -13.53%; 13: -12.49%
| Z-Score (Absolute) |  | 01: 1.02; 02: 0.4; 03: 0; 04: 0.38; 05: 0.19; 06: 0.11; 09: 1.18; 10: 0.67; 11: 2.54; 12: 3; 13: 2.77
| Analytical results for Sn in OREAS 554 (Certified Value 2.22 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 2.3 | 2.2 | 2 | 2.3 | 2.3 | 2 | 1.3 | 2 | < 10 | 2.2 | 2.3 | 2.1 | 2.4 | 2.2 | 2.3 | 2.0211 | 2 | 2.4 | 2.2435 | 2.3 | 2.2 | 2.4 | 2.2 | 1.88 |  |  | 1
| 1 | 2 | 2.2 | 2.4 | 2 | 2.3 | 2.2 | 2 | 1.4 | 2 | < 10 | 2.2 | 2.4 | 2.1 | 2.4 | 2.2 | 2.3 | 1.9952 | 2 | 2.4 | 2.3071 | 2.3 | 2 | 2.5 | 2.1 | 1.96 |  |  | 0
| 1 | 3 | 2.2 | 2.4 | 2 | 2.3 | 2.3 | 2 | 1.4 | 2 | < 10 | 2.2 | 2.3 | 2.1 | 2.3 | 2.2 | 2.4 | 4.1535 | 2 | 2.5 | 2.2607 | 2.3 | 2.1 | 2.5 | 2.1 | 1.93 |  |  | 16
| 1 | 4 | 2.1 | 2.5 | 2 | 2.2 | 2.3 | 2.5 | 1 | 2 | < 10 | 2.1 | 2.4 | 2 | 2.3 | 2.2 | 2.4 | 2.0476 | 2 | 2.4 | 2.2413 | 2.2 | 2 | 2.5 | 2.2 | 1.92 |  |  | 2.2152
| 1 | 5 | 2.2 | 2.1 | 2 | 2.2 | 2.3 | 2 | 1.3 | 2 | < 10 | 2.1 | 2.3 | 2.1 | 2.3 | 2.3 | 2.3 | 1.8452 | 2 | 2.5 | 2.2363 | 2.3 | 1.9 | 2.5 | 2.2 | 2.02 |  |  | 126
| 1 | 6 | 2.2 | 2.3 | 2 | 2.3 | 2.2 | 2 | 1.3 | 2 | < 10 | 2.1 | 2.3 | 2.1 | 2.3 | 2.1 | 2.4 | 2.0103 | 2 | 2.4 | 2.2168 | 2.3 | 2.2 | 2.5 | 2.1 | 1.91
| Mean |  | 2.2 | 2.3167 | 2 | 2.2667 | 2.2667 | 2.0833 | 1.2833 | 2 |  | 2.15 | 2.3333 | 2.0833 | 2.3333 | 2.2 | 2.35 | 2.3455 | 2 | 2.4333 | 2.251 | 2.2833 | 2.0667 | 2.4833 | 2.15 | 1.9367
| Median |  | 2.2 | 2.35 | 2 | 2.3 | 2.3 | 2 | 1.3 | 2 |  | 2.15 | 2.3 | 2.1 | 2.3 | 2.2 | 2.35 | 2.0157 | 2 | 2.4 | 2.2424 | 2.3 | 2.05 | 2.5 | 2.15 | 1.925
| Std Dev. |  | 0.063246 | 0.1472 | 0 | 0.05164 | 0.05164 | 0.20412 | 0.1472 | 0 |  | 0.054772 | 0.05164 | 0.040825 | 0.05164 | 0.063246 | 0.054772 | 0.88861 | 0 | 0.05164 | 0.030911 | 0.040825 | 0.12111 | 0.040825 | 0.054772 | 0.048442
| Rel.Std.Dev. |  | 2.875% | 6.354% | 0% | 2.278% | 2.278% | 9.798% | 11.47% | 0% |  | 2.548% | 2.213% | 1.96% | 2.213% | 2.875% | 2.331% | 37.89% | 0% | 2.122% | 1.373% | 1.788% | 5.86% | 1.644% | 2.548% | 2.501%
| PDM3 |  | -0.6847% | 4.582% | -9.713% | 2.325% | 2.325% | -5.951% | -42.07% | -9.713% |  | -2.942% | 5.334% | -5.951% | 5.334% | -0.6847% | 6.087% | 5.883% | -9.713% | 9.849% | 1.616% | 3.077% | -6.704% | 12.11% | -2.942% | -12.57%
| Z-Score (Absolute) |  | 0.54 | 0.4 | NA | 0 | 0 | 1.48 | 7.96 | NA | 22.12 | 0.94 | 0.54 | 1.48 | 0.54 | 0.54 | 0.67 | 0.64 | NA | 1.35 | 0.13 | 0.13 | 1.62 | 1.75 | 0.94 | 2.67
Results from laboratories 3, 10 & 20 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 554 (Certified Value 67 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 66 | 70.52 | 67 | 65 | 65.4 | 65 | 65 | 64.7 | 67.8 | 60 | 71.9 | 72.9 | 71.2 | 70.6 | 72.2 | 80 | 79.3893 | 66.9 | 66.3 | 65.7548 | 64.1 | 66 | 67.7 | 62.8 | 65 |  | 1
| 1 | 2 | 66.1 | 71.1 | 66 | 65.5 | 64.6 | 70 | 65 | 64.9 | 67.9 | 60 | 72 | 72.5 | 72.2 | 70.9 | 72.8 | 80 | 79.2716 | 66.8 | 65.5 | 67.357 | 64.3 | 66 | 67.2 | 63.2 | 64 |  | 0
| 1 | 3 | 65.7 | 70.85 | 66 | 66 | 64.7 | 65 | 67 | 69.6 | 67.3 | 59 | 71.6 | 71.2 | 70.1 | 69.4 | 71 | 83 | 79.3832 | 67.3 | 67.3 | 64.6921 | 64.1 | 66 | 67.5 | 62.6 | 66 |  | 16
| 1 | 4 | 66.1 | 72.2 | 65.5 | 66 | 65.3 | 65 | 65 | 65 | 68.3 | 58 | 70.9 | 70.5 | 71.1 | 72.1 | 69.8 | 80 | 76.9465 | 65.7 | 64.2 | 66.466 | 63.1 | 68 | 66.9 | 64.4 | 64 |  | 67.204
| 1 | 5 | 66 | 72.18 | 69 | 66 | 66.1 | 65 | 66 | 66.6 | 69.3 | 59 | 67.7 | 72.3 | 70.7 | 69.8 | 70.8 | 81 | 76.9888 | 67.2 | 66.4 | 65.3345 | 62.8 | 64 | 67.5 | 62.1 | 65 |  | 127
| 1 | 6 | 66.2 | 69.87 | 68.5 | 67 | 66.3 | 65 | 65 | 66.9 | 67 | 57 | 68.8 | 68 | 71 | 72 | 68.8 | 81 | 78.3697 | 67.8 | 63.6 | 64.5077 | 64 | 68 | 68.1 | 64.6 | 65
| Mean |  | 66.0167 | 71.12 | 67 | 65.9167 | 65.4 | 65.8333 | 65.5 | 66.2833 | 67.9333 | 58.8333 | 70.4833 | 71.2333 | 71.05 | 70.8 | 70.9 | 80.8333 | 78.3915 | 66.95 | 65.55 | 65.6853 | 63.7333 | 66.3333 | 67.4833 | 63.2833 | 64.8333
| Median |  | 66.05 | 70.975 | 66.5 | 66 | 65.35 | 65 | 65 | 65.8 | 67.85 | 59 | 71.25 | 71.75 | 71.05 | 70.75 | 70.9 | 80.5 | 78.8206 | 67.05 | 65.9 | 65.5446 | 64.05 | 66 | 67.5 | 63 | 65
| Std Dev. |  | 0.17224 | 0.92583 | 1.4491 | 0.66458 | 0.69857 | 2.0412 | 0.83666 | 1.8734 | 0.81158 | 1.169 | 1.806 | 1.8173 | 0.6892 | 1.1082 | 1.4792 | 1.169 | 1.167 | 0.7064 | 1.4124 | 1.0868 | 0.62183 | 1.5055 | 0.41191 | 1.0088 | 0.75277
| Rel.Std.Dev. |  | 0.2609% | 1.302% | 2.163% | 1.008% | 1.068% | 3.101% | 1.277% | 2.826% | 1.195% | 1.987% | 2.562% | 2.551% | 0.97% | 1.565% | 2.086% | 1.446% | 1.489% | 1.055% | 2.155% | 1.654% | 0.9757% | 2.27% | 0.6104% | 1.594% | 1.161%
| PDM3 |  | -1.767% | 5.827% | -0.3036% | -1.916% | -2.684% | -2.04% | -2.536% | -1.37% | 1.085% | -12.46% | 4.88% | 5.996% | 5.723% | 5.351% | 5.5% | 20.28% | 16.65% | -0.378% | -2.461% | -2.26% | -5.164% | -1.296% | 0.4156% | -5.834% | -3.528%
| Z-Score (Absolute) |  | 0.14 | 2.15 | 0.3 | 0.19 | 0.42 | 0.22 | 0.37 | 0.02 | 0.72 | 3.37 | 1.87 | 2.2 | 2.12 | 2.01 | 2.05 | 6.52 | 5.42 | 0.28 | 0.35 | 0.29 | 1.17 | 0 | 0.52 | 1.37 | 0.67
| Analytical results for Ta in OREAS 554 (Certified Value 0.35 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.25 | 0.23 | 0.4 | 0.5 | 0.26 | < 0.5 | < 0.5 | 0.6 | < 5 | 0.32 | 0.48 | 0.43 | 0.42 | 0.38 | NA | 0.26649 | 0.48 | 0.61773 | 0.19 | 0.2 | 0.47 | 0.52 | 0.66 |  |  |  | 1
| 1 | 2 | 0.23 | 0.33 | 0.5 | 0.4 | 0.25 | < 0.5 | < 0.5 | 0.6 | < 5 | 0.32 | 0.51 | 0.43 | 0.41 | 0.37 | NA | 0.22306 | 0.5 | 0.64992 | 0.22 | 0.3 | 0.46 | 0.51 | 0.67 |  |  |  | 0
| 1 | 3 | 0.2 | 0.31 | 0.4 | 0.5 | 0.22 | < 0.5 | < 0.5 | 0.7 | < 5 | 0.34 | 0.46 | 0.43 | 0.42 | 0.35 | NA | 0.2145 | 0.53 | 0.64064 | 0.21 | 0.3 | 0.5 | 0.47 | 0.68 |  |  |  | 16
| 1 | 4 | 0.2 | 0.37 | 0.4 | 0.4 | 0.21 | < 0.5 | < 0.5 | 0.6 | < 5 | 0.33 | 0.49 | 0.4 | 0.39 | 0.36 | NA | 0.19646 | 1 | 0.6374 | 0.22 | 0.2 | 0.46 | 0.44 | 0.67 |  |  |  | 0.34833
| 1 | 5 | 0.2 | 0.36 | 0.5 | 0.4 | 0.25 | < 0.5 | < 0.5 | 0.6 | < 5 | 0.36 | 0.45 | 0.43 | 0.41 | 0.36 | NA | 0.16052 | 0.68 | 0.64148 | 0.23 | 0.2 | 0.46 | 0.52 | 0.72 |  |  |  | 128
| 1 | 6 | 0.2 | 0.34 | 0.4 | 0.4 | 0.25 | < 0.5 | < 0.5 | 0.6 | < 5 | 0.37 | 0.3 | 0.4 | 0.43 | 0.37 | NA | 0.20253 | 0.78 | 0.63304 | 0.23 | 0.3 | 0.45 | 0.47 | 0.7
| Mean |  | 0.21333 | 0.32333 | 0.43333 | 0.43333 | 0.24 |  |  | 0.61667 |  | 0.34 | 0.44833 | 0.42 | 0.41333 | 0.365 |  | 0.21059 | 0.66167 | 0.6367 | 0.21667 | 0.25 | 0.46667 | 0.48833 | 0.68333
| Median |  | 0.2 | 0.335 | 0.4 | 0.4 | 0.25 |  |  | 0.6 |  | 0.335 | 0.47 | 0.43 | 0.415 | 0.365 |  | 0.20851 | 0.605 | 0.63902 | 0.22 | 0.25 | 0.46 | 0.49 | 0.675
| Std Dev. |  | 0.021602 | 0.050465 | 0.05164 | 0.05164 | 0.02 |  |  | 0.040825 |  | 0.020976 | 0.075741 | 0.015492 | 0.013663 | 0.010488 |  | 0.034835 | 0.20262 | 0.010834 | 0.015055 | 0.054772 | 0.017512 | 0.033116 | 0.022509
| Rel.Std.Dev. |  | 10.13% | 15.61% | 11.92% | 11.92% | 8.333% |  |  | 6.62% |  | 6.169% | 16.89% | 3.689% | 3.305% | 2.873% |  | 16.54% | 30.62% | 1.702% | 6.949% | 21.91% | 3.753% | 6.781% | 3.294%
| PDM3 |  | -38.75% | -7.175% | 24.4% | 24.4% | -31.1% |  |  | 77.04% |  | -2.391% | 28.71% | 20.58% | 18.66% | 4.787% |  | -39.54% | 89.96% | 82.79% | -37.8% | -28.23% | 33.97% | 40.19% | 96.18%
| Z-Score (Absolute) |  | 0.85 | 0.32 | NA | NA | 0.72 | 0.67 | 0.67 | NA | 10.23 | 0.24 | 0.29 | 0.15 | 0.12 | 0.12 | NA | 0.87 | 1.32 | 1.2 | 0.84 | NA | 0.38 | 0.48 | 1.43
Results from laboratories 3, 4, 10 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 554 (Certified Value 0.52 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 06; 06: 10; 09: 11; 10: 21; 12: 22; 13: 24; 14: 29; 15: 30 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 09: 4A*MS; 10: 4A*OES/MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*MS |  |  |  | 2
3
| 1 | 1 | 01: 0.54; 02: 0.49; 03: 0.54; 04: 0.54; 05: 0.45; 06: 0.5; 09: 0.75; 10: 0.52; 12: 0.53038; 13: 0.5; 14: 0.43; 15: 0.41 |  |  |  | 1
| 1 | 2 | 01: 0.54; 02: 0.48; 03: 0.52; 04: 0.54; 05: 0.5; 06: 0.6; 09: 0.78; 10: 0.57; 12: 0.5715; 13: 0.4; 14: 0.46; 15: 0.4 |  |  |  | 0
| 1 | 3 | 01: 0.54; 02: 0.5; 03: 0.52; 04: 0.54; 05: 0.45; 06: 0.6; 09: 0.82; 10: 0.57; 12: 0.55613; 13: 0.4; 14: 0.45; 15: 0.39 |  |  |  | 16
| 1 | 4 | 01: 0.54; 02: 0.49; 03: 0.54; 04: 0.54; 05: 0.5; 06: 0.6; 09: 0.83; 10: 0.55; 12: 0.56237; 13: 0.4; 14: 0.45; 15: 0.4 |  |  |  | 0.51813
| 1 | 5 | 01: 0.53; 02: 0.5; 03: 0.52; 04: 0.54; 05: 0.5; 06: 0.6; 09: 0.82; 10: 0.56; 12: 0.55335; 13: 0.5; 14: 0.45; 15: 0.38 |  |  |  | 129
| 1 | 6 | 01: 0.53; 02: 0.49; 03: 0.54; 04: 0.54; 05: 0.5; 06: 0.6; 09: 0.74; 10: 0.57; 12: 0.55228; 13: 0.5; 14: 0.43; 15: 0.41
| Mean |  | 01: 0.53667; 02: 0.49167; 03: 0.53; 04: 0.54; 05: 0.48333; 06: 0.58333; 09: 0.79; 10: 0.55667; 12: 0.55434; 13: 0.45; 14: 0.445; 15: 0.39833
| Median |  | 01: 0.54; 02: 0.49; 03: 0.53; 04: 0.54; 05: 0.5; 06: 0.6; 09: 0.8; 10: 0.565; 12: 0.55474; 13: 0.45; 14: 0.45; 15: 0.4
| Std Dev. |  | 01: 0.005164; 02: 0.0075277; 03: 0.010954; 04: 0; 05: 0.02582; 06: 0.040825; 09: 0.038987; 10: 0.019664; 12: 0.013722; 13: 0.054772; 14: 0.012247; 15: 0.01169
| Rel.Std.Dev. |  | 01: 0.9622%; 02: 1.531%; 03: 2.067%; 04: 0%; 05: 5.342%; 06: 6.999%; 09: 4.935%; 10: 3.532%; 12: 2.475%; 13: 12.17%; 14: 2.752%; 15: 2.935%
| PDM3 |  | 01: 3.579%; 02: -5.107%; 03: 2.292%; 04: 4.222%; 05: -6.715%; 06: 12.59%; 09: 52.47%; 10: 7.439%; 12: 6.989%; 13: -13.15%; 14: -14.11%; 15: -23.12%
| Z-Score (Absolute) |  | 01: 0.07; 02: 0.86; 03: 0.07; 04: 0.14; 05: 1.04; 06: NA; 09: 5.33; 10: 0.48; 12: 0.44; 13: NA; 14: 1.83; 15: 2.8
Results from laboratories 10 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 554 (Certified Value 0.049 ppm) | [27 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | < 0.05 | < 0.2 | < 0.2 | < 0.2 | 0.1 | 6 | < 0.1 | < 5 | 0.06 | 0.05 | < 0.05 | 0.06 | < 0.05 | 0.09 | 0.10929 | < 0.05 | < 0.1 | < 0.05 | < 0.05 | 0.1 | [6 empty] | 1
| 1 | 2 | < 0.05 | < 0.2 | < 0.2 | < 0.2 | 0.1 | 3 | < 0.1 | < 5 | 0.05 | 0.06 | < 0.05 | 0.07 | 0.05 | 0.05 | 0.12936 | 0.06 | < 0.1 | 0.06 | < 0.05 | 0.08 | [6 empty] | 0
| 1 | 3 | < 0.05 | < 0.2 | < 0.2 | < 0.2 | 0.1 | 2 | < 0.1 | < 5 | < 0.05 | 0.06 | 0.07 | 0.07 | 0.05 | 0.06 | 0.10747 | 0.05 | 0.1 | 0.05 | < 0.05 | 0.08 | [6 empty] | 16
| 1 | 4 | 0.07 | < 0.2 | < 0.2 | < 0.2 | 0.1 | 2 | < 0.1 | < 5 | 0.05 | 0.07 | 0.06 | 0.08 | 0.08 | 0.03 | 0.15344 | 0.08 | < 0.1 | < 0.05 | 0.05 | 0.06 | [6 empty] | 0.049133
| 1 | 5 | 0.05 | < 0.2 | < 0.2 | < 0.2 | 0.1 | 2 | < 0.1 | < 5 | 0.05 | 0.06 | < 0.05 | 0.06 | < 0.05 | 0.05 | 0.11058 | 0.07 | < 0.1 | < 0.05 | < 0.05 | 0.09 | [6 empty] | 130
| 1 | 6 | 0.08 | < 0.2 | < 0.2 | < 0.2 | < 0.1 | 1 | < 0.1 | < 5 | < 0.05 | 0.06 | 0.06 | 0.08 | 0.06 | 0.05 | 0.10244 | 0.06 | < 0.1 | 0.08 | 0.08 | 0.09
| Mean |  | 0.066667 |  |  |  | 0.1 | 2.6667 |  |  | 0.0525 | 0.06 | 0.063333 | 0.07 | 0.06 | 0.055 | 0.11876 | 0.064 | 0.1 | 0.063333 | 0.065 | 0.083333
| Median |  | 0.07 |  |  |  | 0.1 | 2 |  |  | 0.05 | 0.06 | 0.06 | 0.07 | 0.055 | 0.05 | 0.10994 | 0.06 | 0.1 | 0.06 | 0.065 | 0.085
| Std Dev. |  | 0.015275 |  |  |  | 0 | 1.7512 |  |  | 0.005 | 0.0063246 | 0.0057735 | 0.0089443 | 0.014142 | 0.019748 | 0.019314 | 0.011402 |  | 0.015275 | 0.021213 | 0.013663
| Rel.Std.Dev. |  | 22.91% |  |  |  | 0% | 65.67% |  |  | 9.524% | 10.54% | 9.116% | 12.78% | 23.57% | 35.91% | 16.26% | 17.82% |  | 24.12% | 32.64% | 16.4%
| PDM3 |  | 35.69% |  |  |  | 103.5% | 5327% |  |  | 6.852% | 22.12% | 28.9% | 42.47% | 22.12% | 11.94% | 141.7% | 30.26% | 103.5% | 28.9% | 32.29% | 69.61%
| Z-Score (Absolute) |  | 0.59 | 2.15 | 2.15 | 2.15 | NA | NA | 0.38 | 123.52 | 0.72 | 0.13 | 0.67 | 0.63 | 0.46 | 0.13 | 3.1 | 0 | NA | 0.67 | 0.97 | 1.31
Results from laboratories 5, 6 & 23 were removed due to their 0.1 ppm or greater reading resolution.
| Analytical results for Th in OREAS 554 (Certified Value 8.99 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
3
| 1 | 1 | 9.4 | 8.81 | 8.4 | 8.5 | 9.17 | 9 | 9.3 | 9 | 10.2 | 8.74 | 9.79 | 9.08 | 9.14 | 9.14 | 9.3 | 9.9619 | 9.3 | 7.7 | 8.6867 | 8.5 | 9.6 | 9.5 | 7.7 | 8.26 |  |  | 1
| 1 | 2 | 9.4 | 8.79 | 8.4 | 7.7 | 9.07 | 8.7 | 9.2 | 8.8 | 10.1 | 9.09 | 9.96 | 9.31 | 9.27 | 9.32 | 9.1 | 10.1939 | 9.1 | 8 | 8.7669 | 8.64 | 10 | 9.5 | 7.8 | 8.55 |  |  | 0
| 1 | 3 | 9.4 | 8.83 | 8.4 | 8.2 | 9.06 | 8.7 | 9.4 | 9.4 | 10.2 | 8.98 | 9.86 | 9.31 | 9.36 | 8.94 | 9.1 | 10.4287 | 9.2 | 8.1 | 8.6304 | 8.44 | 9.2 | 9.6 | 7.7 | 8.51 |  |  | 16
| 1 | 4 | 9.4 | 8.85 | 8.6 | 8.2 | 9.24 | 9.2 | 8.9 | 8.4 | 10.4 | 8.8 | 9.38 | 9.39 | 9.3 | 8.71 | 8.8 | 10.0506 | 9.2 | 7.9 | 8.5179 | 8.29 | 9.3 | 9.7 | 7.8 | 8.44 |  |  | 8.9941
| 1 | 5 | 9.2 | 8.76 | 8.5 | 8.4 | 9.17 | 9.2 | 8.9 | 8.6 | 10.2 | 8.75 | 9.98 | 9.49 | 9.23 | 8.68 | 9.1 | 9.5515 | 9.3 | 8 | 8.5918 | 8.36 | 10 | 9.5 | 7.9 | 8.41 |  |  | 131
| 1 | 6 | 9.3 | 8.69 | 8.7 | 7.8 | 9.19 | 8.7 | 9.2 | 9 | 10.1 | 8.85 | 9.27 | 9.52 | 9.4 | 8.91 | 8.9 | 9.8036 | 9.4 | 7.8 | 8.7558 | 8.55 | 9.1 | 9.5 | 7.7 | 8.58
| Mean |  | 9.35 | 8.7883 | 8.5 | 8.1333 | 9.15 | 8.9167 | 9.15 | 8.8667 | 10.2 | 8.8683 | 9.7067 | 9.35 | 9.2833 | 8.95 | 9.05 | 9.9984 | 9.25 | 7.9167 | 8.6582 | 8.4633 | 9.5333 | 9.55 | 7.7667 | 8.4583
| Median |  | 9.4 | 8.8 | 8.45 | 8.2 | 9.17 | 8.85 | 9.2 | 8.9 | 10.2 | 8.825 | 9.825 | 9.35 | 9.285 | 8.925 | 9.1 | 10.0063 | 9.25 | 7.95 | 8.6585 | 8.47 | 9.45 | 9.5 | 7.75 | 8.475
| Std Dev. |  | 0.083666 | 0.057417 | 0.12649 | 0.32042 | 0.070711 | 0.24833 | 0.20736 | 0.35024 | 0.10954 | 0.13934 | 0.30553 | 0.15887 | 0.093095 | 0.24706 | 0.17607 | 0.30499 | 0.10488 | 0.1472 | 0.09698 | 0.12754 | 0.39833 | 0.083666 | 0.08165 | 0.11652
| Rel.Std.Dev. |  | 0.8948% | 0.6533% | 1.488% | 3.94% | 0.7728% | 2.785% | 2.266% | 3.95% | 1.074% | 1.571% | 3.148% | 1.699% | 1.003% | 2.76% | 1.946% | 3.05% | 1.134% | 1.859% | 1.12% | 1.507% | 4.178% | 0.8761% | 1.051% | 1.378%
| PDM3 |  | 3.957% | -2.288% | -5.494% | -9.57% | 1.733% | -0.8609% | 1.733% | -1.417% | 13.41% | -1.398% | 7.923% | 3.957% | 3.216% | -0.4903% | 0.6216% | 11.17% | 2.845% | -11.98% | -3.734% | -5.901% | 5.995% | 6.181% | -13.65% | -5.957%
| Z-Score (Absolute) |  | 0.67 | 0.41 | 0.96 | 1.67 | 0.29 | 0.16 | 0.29 | 0.26 | 2.31 | 0.25 | 1.36 | 0.67 | 0.55 | 0.1 | 0.1 | 1.92 | 0.48 | 2.09 | 0.66 | 1.03 | 1.03 | 1.06 | 2.38 | 1.04
| Analytical results for Ti in OREAS 554 (Certified Value 0.161 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 3
3
| 1 | 1 | 0.13 | 0.1284 | 0.21 | 0.2 | 0.1546 | 0.134 | 0.136 | 0.219 | 0.1 | 0.08 | 0.15 | 0.187 | 0.178 | 0.167 | 0.155 | 0.1291 | 0.15363 | 0.19 | 0.22317 | 0.111 | 0.1387 | 0.17 | 0.11 | 0.211 |  |  | 1
| 1 | 2 | 0.13 | 0.1499 | 0.2 | 0.19 | 0.1499 | 0.138 | 0.137 | 0.223 | 0.1 | 0.08 | 0.152 | 0.188 | 0.178 | 0.166 | 0.158 | 0.1312 | 0.14053 | 0.19 | 0.23185 | 0.11 | 0.1455 | 0.17 | 0.12 | 0.2104 |  |  | 0
| 1 | 3 | 0.13 | 0.1468 | 0.2 | 0.185 | 0.137 | 0.132 | 0.14 | 0.225 | 0.1 | 0.07 | 0.155 | 0.181 | 0.18 | 0.166 | 0.15 | 0.1318 | 0.1359 | 0.2 | 0.23222 | 0.115 | 0.1387 | 0.17 | 0.12 | 0.2134 |  |  | 16
| 1 | 4 | 0.13 | 0.1662 | 0.2 | 0.2 | 0.1412 | 0.132 | 0.141 | 0.224 | 0.11 | 0.07 | 0.15 | 0.188 | 0.174 | 0.161 | 0.15 | 0.1312 | 0.13972 | 0.19 | 0.2254 | 0.118 | 0.1465 | 0.18 | 0.13 | 0.2128 |  |  | 0.1609
| 1 | 5 | 0.12 | 0.1579 | 0.2 | 0.195 | 0.1534 | 0.13 | 0.141 | 0.221 | 0.1 | 0.07 | 0.157 | 0.176 | 0.181 | 0.165 | 0.156 | 0.1301 | 0.13595 | 0.19 | 0.22474 | 0.115 | 0.1339 | 0.17 | 0.14 | 0.214 |  |  | 132
| 1 | 6 | 0.12 | 0.1553 | 0.19 | 0.19 | 0.1604 | 0.135 | 0.136 | 0.215 | 0.1 | 0.07 | 0.156 | 0.149 | 0.177 | 0.171 | 0.154 | 0.1287 | 0.13959 | 0.19 | 0.2205 | 0.113 | 0.1504 | 0.17 | 0.12 | 0.2152
| Mean |  | 0.12667 | 0.15075 | 0.2 | 0.19333 | 0.14942 | 0.1335 | 0.1385 | 0.22117 | 0.10167 | 0.073333 | 0.15333 | 0.17817 | 0.178 | 0.166 | 0.15383 | 0.13035 | 0.14089 | 0.19167 | 0.22631 | 0.11367 | 0.14228 | 0.17167 | 0.12333 | 0.2128
| Median |  | 0.13 | 0.1526 | 0.2 | 0.1925 | 0.15165 | 0.133 | 0.1385 | 0.222 | 0.1 | 0.07 | 0.1535 | 0.184 | 0.178 | 0.166 | 0.1545 | 0.13065 | 0.13966 | 0.19 | 0.22507 | 0.114 | 0.1421 | 0.17 | 0.12 | 0.2131
| Std Dev. |  | 0.005164 | 0.012856 | 0.0063246 | 0.0060553 | 0.0087782 | 0.0028107 | 0.002429 | 0.0037103 | 0.0040825 | 0.005164 | 0.0030768 | 0.015065 | 0.0024495 | 0.0032249 | 0.0032506 | 0.0012566 | 0.0065549 | 0.0040825 | 0.0047439 | 0.0029439 | 0.0061639 | 0.0040825 | 0.010328 | 0.0018199
| Rel.Std.Dev. |  | 4.077% | 8.528% | 3.162% | 3.132% | 5.875% | 2.105% | 1.754% | 1.678% | 4.016% | 7.042% | 2.007% | 8.456% | 1.376% | 1.943% | 2.113% | 0.964% | 4.653% | 2.13% | 2.096% | 2.59% | 4.332% | 2.378% | 8.374% | 0.8552%
| PDM3 |  | -21.27% | -6.306% | 24.3% | 20.16% | -7.134% | -17.03% | -13.92% | 37.46% | -36.81% | -54.42% | -4.7% | 10.73% | 10.63% | 3.172% | -4.389% | -18.98% | -12.44% | 19.12% | 40.66% | -29.35% | -11.57% | 6.694% | -23.35% | 32.26%
| Z-Score (Absolute) |  | 0.67 | 0.03 | 1.26 | 1.08 | 0.07 | 0.49 | 0.36 | 1.82 | 1.32 | 2.07 | 0.03 | 0.69 | 0.68 | 0.37 | 0.05 | 0.57 | 0.29 | 1.04 | 1.95 | 1.01 | 0.26 | 0.52 | 0.75 | 1.6
| Analytical results for Tl in OREAS 554 (Certified Value 0.53 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 0.52 | 0.54 | 0.5 | 0.3 | 0.52 | 0.6 | 0.53 | 0.6 | 0.55 | 0.54 | 0.44 | 0.46 | 0.59 | 0.478 | 0.57803 | < 5 | 0.51 | 0.51411 | 0.54 | 0.57 | 0.52 | 0.44 |  |  |  |  | 1
| 1 | 2 | 0.5 | 0.56 | 0.5 | 0.3 | 0.51 | 0.6 | 0.54 | 0.6 | 0.53 | 0.53 | 0.47 | 0.49 | 0.56 | 0.481 | 0.63141 | < 5 | 0.52 | 0.54729 | 0.52 | 0.51 | 0.52 | 0.46 |  |  |  |  | 0
| 1 | 3 | 0.52 | 0.53 | 0.5 | 0.3 | 0.52 | 0.5 | 0.59 | 0.5 | 0.53 | 0.51 | 0.46 | 0.47 | 0.63 | 0.483 | 0.56622 | < 5 | 0.52 | 0.53609 | 0.53 | 0.52 | 0.52 | 0.45 |  |  |  |  | 16
| 1 | 4 | 0.51 | 0.56 | 0.5 | 0.3 | 0.52 | 0.6 | 0.55 | 0.6 | 0.55 | 0.52 | 0.46 | 0.46 | 0.57 | 0.468 | 0.55291 | < 5 | 0.49 | 0.52966 | 0.5 | 0.53 | 0.51 | 0.44 |  |  |  |  | 0.52737
| 1 | 5 | 0.51 | 0.55 | 0.5 | 0.3 | 0.53 | 0.55 | 0.55 | 0.6 | 0.51 | 0.54 | 0.47 | 0.47 | 0.5 | 0.477 | 0.54935 | < 5 | 0.52 | 0.53329 | 0.53 | 0.55 | 0.53 | 0.46 |  |  |  |  | 133
| 1 | 6 | 0.5 | 0.54 | 0.5 | 0.3 | 0.53 | 0.55 | 0.55 | 0.6 | 0.5 | 0.53 | 0.44 | 0.49 | 0.56 | 0.477 | 0.56933 | < 5 | 0.5 | 0.53374 | 0.54 | 0.52 | 0.52 | 0.44
| Mean |  | 0.51 | 0.54667 | 0.5 | 0.3 | 0.52167 | 0.56667 | 0.55167 | 0.58333 | 0.52833 | 0.52833 | 0.45667 | 0.47333 | 0.56833 | 0.47733 | 0.57454 |  | 0.51 | 0.53236 | 0.52667 | 0.53333 | 0.52 | 0.44833
| Median |  | 0.51 | 0.545 | 0.5 | 0.3 | 0.52 | 0.575 | 0.55 | 0.6 | 0.53 | 0.53 | 0.46 | 0.47 | 0.565 | 0.4775 | 0.56778 |  | 0.515 | 0.53351 | 0.53 | 0.525 | 0.52 | 0.445
| Std Dev. |  | 0.0089443 | 0.012111 | 0 | 0 | 0.0075277 | 0.040825 | 0.020412 | 0.040825 | 0.020412 | 0.01169 | 0.013663 | 0.013663 | 0.042622 | 0.005164 | 0.029816 |  | 0.012649 | 0.01077 | 0.015055 | 0.022509 | 0.0063246 | 0.0098319
| Rel.Std.Dev. |  | 1.754% | 2.215% | 0% | 0% | 1.443% | 7.204% | 3.7% | 6.999% | 3.864% | 2.213% | 2.992% | 2.886% | 7.5% | 1.082% | 5.189% |  | 2.48% | 2.023% | 2.859% | 4.22% | 1.216% | 2.193%
| PDM3 |  | -3.293% | 3.66% | -5.189% | -43.11% | -1.081% | 7.452% | 4.608% | 10.61% | 0.1833% | 0.1833% | -13.41% | -10.25% | 7.768% | -9.487% | 8.946% |  | -3.293% | 0.9475% | -0.1327% | 1.131% | -1.397% | -14.99%
| Z-Score (Absolute) |  | 0.54 | 0.6 | NA | 7.08 | 0.18 | 1.22 | 0.75 | NA | 0.03 | 0.03 | 2.2 | 1.69 | 1.27 | 1.56 | 1.46 | 61.39 | 0.54 | 0.15 | 0.03 | 0.18 | 0.23 | 2.46
Results from laboratories 3 & 11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 554 (Certified Value 0.2 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 10: Lab; 11: Lab; 13: Lab |  |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 06; 05: 10; 06: 11; 10: 22; 11: 24; 13: 29 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*OES/MS; 06: 4A*MS; 10: 4A*MS; 11: 4A*MS; 13: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 01: 0.19; 02: 0.19; 03: 0.22; 04: 0.2; 05: 0.2; 06: 0.21; 10: 0.27124; 11: 0.2; 13: 0.19 |  |  |  |  | 1
| 1 | 2 | 01: 0.19; 02: 0.2; 03: 0.2; 04: 0.2; 05: 0.2; 06: 0.22; 10: 0.26429; 11: 0.2; 13: 0.19 |  |  |  |  | 0
| 1 | 3 | 01: 0.19; 02: 0.21; 03: 0.2; 04: 0.2; 05: 0.3; 06: 0.21; 10: 0.28128; 11: 0.2; 13: 0.19 |  |  |  |  | 16
| 1 | 4 | 01: 0.19; 02: 0.21; 03: 0.22; 04: 0.2; 05: 0.3; 06: 0.23; 10: 0.27045; 11: 0.2; 13: 0.18 |  |  |  |  | 0.20233
| 1 | 5 | 01: 0.18; 02: 0.2; 03: 0.22; 04: 0.2; 05: 0.3; 06: 0.23; 10: 0.2728; 11: 0.2; 13: 0.19 |  |  |  |  | 134
| 1 | 6 | 01: 0.19; 02: 0.22; 03: 0.2; 04: 0.2; 05: 0.2; 06: 0.23; 10: 0.28355; 11: 0.2; 13: 0.18
| Mean |  | 01: 0.18833; 02: 0.205; 03: 0.21; 04: 0.2; 05: 0.25; 06: 0.22167; 10: 0.27393; 11: 0.2; 13: 0.18667
| Median |  | 01: 0.19; 02: 0.205; 03: 0.21; 04: 0.2; 05: 0.25; 06: 0.225; 10: 0.27202; 11: 0.2; 13: 0.19
| Std Dev. |  | 01: 0.0040825; 02: 0.010488; 03: 0.010954; 04: 3.0405e-17; 05: 0.054772; 06: 0.0098319; 10: 0.0072128; 11: 3.0405e-17; 13: 0.005164
| Rel.Std.Dev. |  | 01: 2.168%; 02: 5.116%; 03: 5.216%; 04: 1.52e-14%; 05: 21.91%; 06: 4.435%; 10: 2.633%; 11: 1.52e-14%; 13: 2.766%
| PDM3 |  | 01: -6.919%; 02: 1.318%; 03: 3.789%; 04: -1.153%; 05: 23.56%; 06: 9.555%; 10: 35.39%; 11: -1.153%; 13: -7.743%
| Z-Score (Absolute) |  | 01: 0.78; 02: 0.1; 03: 0.1; 04: NA; 05: NA; 06: 0.57; 10: 2.69; 11: NA; 13: 0.84
Results from laboratories 6, 10 & 24 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 554 (Certified Value 13.8 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.18 | 14.24 | 14.2 | 14 | 14.18 | 14.9 | 13.1 | 14.8 | 15.9 | 14.5 | 14.7 | 15.8 | 13.9 | 13.8 | 12.9 | 12.3704 | 13.2 | 13.3687 | 13.3 | 14.4 | 15.2 | 12.1 | 11.6 |  |  |  | 1
| 1 | 2 | 14 | 14.06 | 14.1 | 14.1 | 14.04 | 14.5 | 12.8 | 13.9 | 16 | 14.1 | 14.8 | 16.2 | 14.1 | 13.8 | 12.9 | 12.5086 | 13.2 | 13.6037 | 13.5 | 13.5 | 14.9 | 12.5 | 12.3 |  |  |  | 0
| 1 | 3 | 14.02 | 14.18 | 13.8 | 13.8 | 14.04 | 14.7 | 13.4 | 15.2 | 16.1 | 14.7 | 14.3 | 17.6 | 13.8 | 14.2 | 12.7 | 12.4637 | 13.5 | 13.4328 | 13.6 | 13.2 | 14.8 | 12.4 | 11.7 |  |  |  | 16
| 1 | 4 | 14.18 | 14.3 | 13.6 | 13.4 | 14 | 14.9 | 12.9 | 14.3 | 16.1 | 13.8 | 14.3 | 16.1 | 13.9 | 14.6 | 12.4 | 11.8208 | 13.1 | 13.3049 | 13.2 | 13.2 | 14.9 | 12.6 | 12 |  |  |  | 13.7747
| 1 | 5 | 14.08 | 14 | 13.7 | 13.5 | 14.2 | 14.8 | 12.7 | 15.2 | 16.5 | 13.4 | 15.2 | 17 | 14.4 | 13.4 | 12.8 | 11.9182 | 13 | 13.8462 | 13.7 | 14.5 | 14.9 | 12.6 | 11.6 |  |  |  | 135
| 1 | 6 | 14.37 | 14.03 | 14.4 | 13.9 | 14.15 | 14.2 | 13.3 | 14 | 15.7 | 13.5 | 13.9 | 16.3 | 15 | 13.2 | 12.6 | 12.142 | 13 | 13.5744 | 13.5 | 13.9 | 15 | 12.6 | 12
| Mean |  | 14.1383 | 14.135 | 13.9667 | 13.7833 | 14.1017 | 14.6667 | 13.0333 | 14.5667 | 16.05 | 14 | 14.5333 | 16.5 | 14.1833 | 13.8333 | 12.7167 | 12.204 | 13.1667 | 13.5218 | 13.4667 | 13.7833 | 14.95 | 12.4667 | 11.8667
| Median |  | 14.13 | 14.12 | 13.95 | 13.85 | 14.095 | 14.75 | 13 | 14.55 | 16.05 | 13.95 | 14.5 | 16.25 | 14 | 13.8 | 12.75 | 12.2562 | 13.15 | 13.5036 | 13.5 | 13.7 | 14.9 | 12.55 | 11.85
| Std Dev. |  | 0.13688 | 0.1226 | 0.31411 | 0.27869 | 0.084951 | 0.27325 | 0.28048 | 0.58195 | 0.26646 | 0.52915 | 0.45898 | 0.66933 | 0.4535 | 0.51251 | 0.19408 | 0.28993 | 0.18619 | 0.19654 | 0.18619 | 0.57764 | 0.13784 | 0.19664 | 0.28048
| Rel.Std.Dev. |  | 0.9682% | 0.8673% | 2.249% | 2.022% | 0.6024% | 1.863% | 2.152% | 3.995% | 1.66% | 3.78% | 3.158% | 4.057% | 3.197% | 3.705% | 1.526% | 2.376% | 1.414% | 1.453% | 1.383% | 4.191% | 0.922% | 1.577% | 2.364%
| PDM3 |  | 2.64% | 2.615% | 1.393% | 0.06245% | 2.373% | 6.475% | -5.382% | 5.749% | 16.52% | 1.635% | 5.507% | 19.78% | 2.966% | 0.4254% | -7.681% | -11.4% | -4.414% | -1.836% | -2.236% | 0.06245% | 8.532% | -9.496% | -13.85%
| Z-Score (Absolute) |  | 0.2 | 0.2 | 0 | 0.22 | 0.16 | 0.83 | 1.11 | 0.71 | 2.48 | 0.04 | 0.67 | 3.01 | 0.26 | 0.16 | 1.49 | 2.1 | 0.95 | 0.53 | 0.59 | 0.22 | 1.17 | 1.78 | 2.5
| Analytical results for V in OREAS 554 (Certified Value 112 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 112 | 113 | 115 | 110 | 120 | 110 | 113 | 113 | 88 | 114 | 116 | 112 | 110 | 112 | 95 | 107.723 | 109 | 120.9482 | 120 | 105 | 112 | 105 | 116 |  |  |  | 1
| 1 | 2 | 113 | 117 | 110 | 105 | 114 | 110 | 116 | 111 | 89 | 118 | 114 | 113 | 106 | 110 | 95 | 113.63 | 106 | 122.9731 | 118 | 108 | 112 | 104 | 113 |  |  |  | 0
| 1 | 3 | 112 | 115 | 105 | 110 | 116 | 110 | 114 | 121 | 88 | 118 | 110 | 111 | 107 | 108 | 97 | 109.1 | 111 | 124.8816 | 115 | 107 | 111 | 109 | 118 |  |  |  | 16
| 1 | 4 | 114 | 116 | 110 | 110 | 115 | 110 | 112 | 116 | 87 | 115 | 113 | 111 | 108 | 106 | 94 | 106.414 | 108 | 121.8101 | 115 | 110 | 113 | 107 | 113 |  |  |  | 112.2119
| 1 | 5 | 113 | 114 | 110 | 105 | 121 | 110 | 113 | 115 | 89 | 110 | 112 | 113 | 109 | 110 | 96 | 102.333 | 111 | 120.5275 | 118 | 107 | 112 | 113 | 116 |  |  |  | 136
| 1 | 6 | 112 | 117 | 110 | 105 | 119 | 110 | 113 | 114 | 88 | 111 | 110 | 114 | 111 | 106 | 95 | 103.323 | 110 | 118.0386 | 118 | 112 | 109 | 110 | 119
| Mean |  | 112.6667 | 115.3333 | 110 | 107.5 | 117.5 | 110 | 113.5 | 115 | 88.1667 | 114.3333 | 112.5 | 112.3333 | 108.5 | 108.6667 | 95.3333 | 107.0872 | 109.1667 | 121.5298 | 117.3333 | 108.1667 | 111.5 | 108 | 115.8333
| Median |  | 112.5 | 115.5 | 110 | 107.5 | 117.5 | 110 | 113 | 114.5 | 88 | 114.5 | 112.5 | 112.5 | 108.5 | 109 | 95 | 107.0685 | 109.5 | 121.3791 | 118 | 107.5 | 112 | 108 | 116
| Std Dev. |  | 0.8165 | 1.633 | 3.1623 | 2.7386 | 2.881 | 0 | 1.3784 | 3.4059 | 0.75277 | 3.3862 | 2.3452 | 1.2111 | 1.8708 | 2.4221 | 1.0328 | 4.1107 | 1.9408 | 2.3207 | 1.9664 | 2.4833 | 1.3784 | 3.3466 | 2.4833
| Rel.Std.Dev. |  | 0.7247% | 1.416% | 2.875% | 2.548% | 2.452% | 0% | 1.214% | 2.962% | 0.8538% | 2.962% | 2.085% | 1.078% | 1.724% | 2.229% | 1.083% | 3.839% | 1.778% | 1.91% | 1.676% | 2.296% | 1.236% | 3.099% | 2.144%
| PDM3 |  | 0.4053% | 2.782% | -1.971% | -4.199% | 4.713% | -1.971% | 1.148% | 2.485% | -21.43% | 1.891% | 0.2567% | 0.1082% | -3.308% | -3.159% | -15.04% | -4.567% | -2.714% | 8.304% | 4.564% | -3.605% | -0.6344% | -3.754% | 3.227%
| Z-Score (Absolute) |  | 0.24 | 0.78 | 0.3 | 0.81 | 1.21 | 0.3 | 0.4 | 0.71 | 4.72 | 0.57 | 0.2 | 0.17 | 0.61 | 0.57 | 3.27 | 0.89 | 0.47 | 2.03 | 1.18 | 0.67 | 0 | 0.71 | 0.88
| Analytical results for W in OREAS 554 (Certified Value 1.27 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 2
2
| 1 | 1 | 1.1 | 1.2 | 1.5 | 1.5 | 1.1 | < 1 | 1 | 1.6 | 1.2 | < 10 | 1.1 | 1.3 | 1.1 | 1.2 | 1.2 | 1.1 | 1.2745 | 1.3 | 1.4056 | 1 | 1.1 | 1.5 | 1.1 | 1.2 |  |  | 1
| 1 | 2 | 1.1 | 1.5 | 1.5 | 1.5 | 1.1 | < 1 | 1 | 1.6 | 1.2 | < 10 | 1.3 | 1.4 | 1.1 | 1.2 | 1.4 | 1.1 | 2.1399 | 1.3 | 1.4748 | 1 | 1.1 | 1.6 | 1.1 | 1.3 |  |  | 0
| 1 | 3 | 1.1 | 1.2 | 1.5 | 1.5 | 1.1 | < 1 | 1.2 | 1.7 | 1.2 | < 10 | 1.2 | 1.3 | 1.2 | 1.4 | 1.3 | 1.1 | 1.4995 | 1.4 | 1.4813 | 1 | 1 | 1.5 | 1.1 | 1.4 |  |  | 16
| 1 | 4 | 1.1 | 1.3 | 1.5 | 1.5 | 1.1 | < 1 | 1.3 | 1.6 | 1.3 | < 10 | 1.1 | 1.4 | 1.1 | 1.3 | 1.2 | 1 | 1.2421 | 1.4 | 1.4105 | 1 | 1.1 | 1.6 | 1.2 | 1.4 |  |  | 1.2697
| 1 | 5 | 1.1 | 1.2 | 1.5 | 1.5 | 1.1 | < 1 | 1.1 | 1.5 | 1.1 | < 10 | 1.3 | 1.3 | 1.8 | 1.2 | 1.2 | 1.1 | 1.2056 | 1.4 | 1.449 | 1 | 1 | 1.7 | 1.3 | 1.5 |  |  | 137
| 1 | 6 | 1.1 | 1.3 | 1.5 | 1 | 1.1 | < 1 | 1.9 | 1.5 | 1.1 | < 10 | 1.3 | 1.2 | 1.1 | 2.7 | 1.3 | 1.1 | 2.102 | 1.4 | 1.4527 | 1.1 | 1.1 | 1.6 | 1.1 | 1.4
| Mean |  | 1.1 | 1.2833 | 1.5 | 1.4167 | 1.1 |  | 1.25 | 1.5833 | 1.1833 |  | 1.2167 | 1.3167 | 1.2333 | 1.5 | 1.2667 | 1.0833 | 1.5773 | 1.3667 | 1.4456 | 1.0167 | 1.0667 | 1.5833 | 1.15 | 1.3667
| Median |  | 1.1 | 1.25 | 1.5 | 1.5 | 1.1 |  | 1.15 | 1.6 | 1.2 |  | 1.25 | 1.3 | 1.1 | 1.25 | 1.25 | 1.1 | 1.387 | 1.4 | 1.4508 | 1 | 1.1 | 1.6 | 1.1 | 1.4
| Std Dev. |  | 2.4324e-16 | 0.1169 | 0 | 0.20412 | 2.4324e-16 |  | 0.33912 | 0.075277 | 0.075277 |  | 0.098319 | 0.075277 | 0.28048 | 0.5933 | 0.08165 | 0.040825 | 0.4336 | 0.05164 | 0.031694 | 0.040825 | 0.05164 | 0.075277 | 0.083666 | 0.10328
| Rel.Std.Dev. |  | 2.211e-14% | 9.109% | 0% | 14.41% | 2.211e-14% |  | 27.13% | 4.754% | 6.361% |  | 8.081% | 5.717% | 22.74% | 39.55% | 6.446% | 3.768% | 27.49% | 3.779% | 2.192% | 4.016% | 4.841% | 4.754% | 7.275% | 7.557%
| PDM3 |  | -13.37% | 1.07% | 18.13% | 11.57% | -13.37% |  | -1.555% | 24.7% | -6.805% |  | -4.18% | 3.695% | -2.868% | 18.13% | -0.2425% | -14.68% | 24.22% | 7.633% | 13.85% | -19.93% | -15.99% | 24.7% | -9.431% | 7.633%
| Z-Score (Absolute) |  | 0.68 | 0.03 | 0.88 | 0.55 | 0.68 | 3.02 | 0.1 | 1.2 | 0.36 | 14.53 | 0.23 | 0.16 | 0.16 | 0.88 | 0.03 | 0.75 | 1.18 | 0.36 | 0.67 | 1.01 | 0.81 | 1.2 | 0.49 | 0.36
| Analytical results for Y in OREAS 554 (Certified Value 12.9 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 12.6 | 12.3 | 14.5 | 13.3 | 12.5 | 12 | 15.4 | 12.5 | 10 | 13.2 | 14.5 | 14.4 | 13.3 | 12.3 | 12.5 | 14.1259 | 14 | 20.0754 | 11.3 | 12.2 | 13.4 | 13.3 | 12.4 |  |  |  | 1
| 1 | 2 | 12.4 | 12.63 | 14.1 | 12.5 | 12.4 | 12.1 | 15 | 12 | 10 | 13.6 | 14.6 | 14.3 | 13.4 | 13.3 | 12.7 | 14.5886 | 14.3 | 20.5146 | 11.1 | 13 | 13.2 | 12.6 | 11.4 |  |  |  | 0
| 1 | 3 | 12.1 | 12.79 | 14.1 | 11.6 | 12.5 | 12.1 | 16.4 | 12 | 9 | 13.7 | 13.8 | 13.4 | 13.2 | 12.3 | 12.9 | 14.9088 | 14.6 | 19.5418 | 11.2 | 12.4 | 12.7 | 12.9 | 11.6 |  |  |  | 16
| 1 | 4 | 12.4 | 12.49 | 14 | 12.1 | 13.2 | 11.9 | 16.6 | 12.8 | 10 | 13.8 | 14.2 | 13.5 | 16.3 | 13 | 13.5 | 14.6078 | 13.8 | 20.0592 | 11.4 | 12.8 | 13.8 | 12.7 | 11.1 |  |  |  | 12.9139
| 1 | 5 | 12 | 12.36 | 14.3 | 12.7 | 13 | 11.9 | 16.4 | 12.8 | 10 | 12.5 | 14.2 | 13.5 | 14.1 | 12.2 | 13 | 13.7915 | 13.9 | 20.2919 | 11.9 | 12.4 | 13 | 11.9 | 11.1 |  |  |  | 138
| 1 | 6 | 12.2 | 12.55 | 14.1 | 13.6 | 12.2 | 12.1 | 15.6 | 11.8 | 10 | 12.7 | 13.2 | 13.6 | 13.5 | 12 | 12.9 | 14.0242 | 13.7 | 20.0943 | 11.7 | 12.9 | 12.7 | 13.2 | 11.3
| Mean |  | 12.2833 | 12.52 | 14.1833 | 12.6333 | 12.6333 | 12.0167 | 15.9 | 12.3167 | 9.8333 | 13.25 | 14.0833 | 13.7833 | 13.9667 | 12.5167 | 12.9167 | 14.3411 | 14.05 | 20.0962 | 11.4333 | 12.6167 | 13.1333 | 12.7667 | 11.4833
| Median |  | 12.3 | 12.52 | 14.1 | 12.6 | 12.5 | 12.05 | 16 | 12.25 | 10 | 13.4 | 14.2 | 13.55 | 13.45 | 12.3 | 12.9 | 14.3573 | 13.95 | 20.0849 | 11.35 | 12.6 | 13.1 | 12.8 | 11.35
| Std Dev. |  | 0.22286 | 0.17933 | 0.18348 | 0.74207 | 0.38297 | 0.098319 | 0.65422 | 0.44008 | 0.40825 | 0.54681 | 0.51543 | 0.4446 | 1.186 | 0.51153 | 0.33714 | 0.42508 | 0.33912 | 0.32329 | 0.30768 | 0.32506 | 0.4274 | 0.50465 | 0.48751
| Rel.Std.Dev. |  | 1.814% | 1.432% | 1.294% | 5.874% | 3.031% | 0.8182% | 4.115% | 3.573% | 4.152% | 4.127% | 3.66% | 3.226% | 8.492% | 4.087% | 2.61% | 2.964% | 2.414% | 1.609% | 2.691% | 2.576% | 3.254% | 3.953% | 4.245%
| PDM3 |  | -4.883% | -3.05% | 9.83% | -2.173% | -2.173% | -6.948% | 23.12% | -4.625% | -23.85% | 2.603% | 9.056% | 6.733% | 8.152% | -3.076% | 0.0215% | 11.05% | 8.798% | 55.62% | -11.46% | -2.302% | 1.699% | -1.14% | -11.08%
| Z-Score (Absolute) |  | 0.54 | 0.31 | 1.31 | 0.2 | 0.2 | 0.8 | 2.98 | 0.51 | NA | 0.4 | 1.21 | 0.92 | 1.1 | 0.32 | 0.07 | 1.46 | 1.18 | 7.07 | 1.37 | 0.22 | 0.28 | 0.07 | 1.32
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 554 (Certified Value 1.42 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 05: 04; 06: 06; 09: 10; 10: 11; 11: 21; 12: 22; 13: 24; 14: 29; 15: 30 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 09: 4A*OES/MS; 10: 4A*MS; 11: 4A*OES/MS; 12: 4A*MS; 13: 4A*MS; 14: 4A*OES/MS; 15: 4A*MS |  |  |  | 2
2
| 1 | 1 | 01: 1.3; 02: 1.29; 03: 1.5; 05: 1.45; 06: 1.5; 09: 1.6; 10: 1.3; 11: 1.5; 12: 1.8576; 13: 1.3; 14: 1.3; 15: 1.05 |  |  |  | 1
| 1 | 2 | 01: 1.3; 02: 1.28; 03: 1.5; 05: 1.4; 06: 1.55; 09: 1.6; 10: 1.4; 11: 1.5; 12: 1.8499; 13: 1.2; 14: 1.2; 15: 1.19 |  |  |  | 0
| 1 | 3 | 01: 1.3; 02: 1.31; 03: 1.55; 05: 1.5; 06: 1.45; 09: 1.7; 10: 1.4; 11: 1.5; 12: 1.8407; 13: 1.2; 14: 1.3; 15: 1.11 |  |  |  | 16
| 1 | 4 | 01: 1.3; 02: 1.3; 03: 1.4; 05: 1.55; 06: 1.5; 09: 1.7; 10: 1.4; 11: 1.5; 12: 1.7769; 13: 1.2; 14: 1.3; 15: 1.12 |  |  |  | 1.4162
| 1 | 5 | 01: 1.3; 02: 1.27; 03: 1.5; 05: 1.35; 06: 1.5; 09: 1.7; 10: 1.4; 11: 1.5; 12: 1.7798; 13: 1.3; 14: 1.3; 15: 1.1 |  |  |  | 139
| 1 | 6 | 01: 1.3; 02: 1.32; 03: 1.4; 05: 1.55; 06: 1.35; 09: 1.7; 10: 1.3; 11: 1.4; 12: 1.8551; 13: 1.2; 14: 1.2; 15: 1.12
| Mean |  | 01: 1.3; 02: 1.295; 03: 1.475; 05: 1.4667; 06: 1.475; 09: 1.6667; 10: 1.3667; 11: 1.4833; 12: 1.8267; 13: 1.2333; 14: 1.2667; 15: 1.115
| Median |  | 01: 1.3; 02: 1.295; 03: 1.5; 05: 1.475; 06: 1.5; 09: 1.7; 10: 1.4; 11: 1.5; 12: 1.8453; 13: 1.2; 14: 1.3; 15: 1.115
| Std Dev. |  | 01: 0; 02: 0.018708; 03: 0.061237; 05: 0.08165; 06: 0.06892; 09: 0.05164; 10: 0.05164; 11: 0.040825; 12: 0.037906; 13: 0.05164; 14: 0.05164; 15: 0.045056
| Rel.Std.Dev. |  | 01: 0%; 02: 1.445%; 03: 4.152%; 05: 5.567%; 06: 4.673%; 09: 3.098%; 10: 3.779%; 11: 2.752%; 12: 2.075%; 13: 4.187%; 14: 4.077%; 15: 4.041%
| PDM3 |  | 01: -8.208%; 02: -8.561%; 03: 4.148%; 05: 3.56%; 06: 4.148%; 09: 17.68%; 10: -3.501%; 11: 4.737%; 12: 28.98%; 13: -12.92%; 14: -10.56%; 15: -21.27%
| Z-Score (Absolute) |  | 01: 0.66; 02: 0.69; 03: 0.33; 05: 0.28; 06: 0.33; 09: 1.41; 10: 0.28; 11: 0.38; 12: 2.32; 13: 1.04; 14: 0.85; 15: 1.71
| Analytical results for Zn in OREAS 554 (Certified Value 55 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
1
| 1 | 1 | 50 | 55 | 52 | 50 | 56 | 55 | 52 | 56.5 | 59.2 | 42 | 55 | 56 | 54 | 54 | 52 | 59 | 42.0085 | 58 | 54.9613 | 85 | 58.2 | 53 | 54 | 84 |  |  | 1
| 1 | 2 | 48 | 56 | 54 | 50 | 53 | 60 | 58 | 53.9 | 58.3 | 44 | 57 | 54 | 57 | 53 | 55 | 57 | 40.5134 | 65 | 53.8073 | 84 | 57.9 | 53 | 54 | 86 |  |  | 0
| 1 | 3 | 51 | 56 | 52 | 50 | 52 | 60 | 53 | 58 | 60.2 | 44 | 55 | 53 | 56 | 52 | 54 | 58 | 39.6928 | 60 | 55.0097 | 83 | 59.3 | 53 | 53 | 85 |  |  | 16
| 1 | 4 | 51 | 55 | 52 | 50 | 55 | 60 | 59 | 55.6 | 59.8 | 44 | 53 | 57 | 55 | 53 | 52 | 56 | 38.5255 | 57 | 55.9342 | 82 | 64.4 | 52 | 56 | 86 |  |  | 55.2939
| 1 | 5 | 50 | 55 | 54 | 50 | 55 | 60 | 60 | 55.7 | 58.5 | 44 | 56 | 57 | 56 | 53 | 54 | 57 | 37.5566 | 59 | 53.8137 | 85 | 58.3 | 55 | 58 | 85 |  |  | 140
| 1 | 6 | 53 | 55 | 54 | 50 | 55 | 60 | 58 | 56.3 | 59.5 | 44 | 55 | 62 | 57 | 55 | 52 | 57 | 40.7657 | 56 | 53.0789 | 86 | 60.6 | 53 | 55 | 85
| Mean |  | 50.5 | 55.3333 | 53 | 50 | 54.3333 | 59.1667 | 56.6667 | 56 | 59.25 | 43.6667 | 55.1667 | 56.5 | 55.8333 | 53.3333 | 53.1667 | 57.3333 | 39.8437 | 59.1667 | 54.4342 | 84.1667 | 59.7833 | 53.1667 | 55 | 85.1667
| Median |  | 50.5 | 55 | 53 | 50 | 55 | 60 | 58 | 56 | 59.35 | 44 | 55 | 56.5 | 56 | 53 | 53 | 57 | 40.1031 | 58.5 | 54.3875 | 84.5 | 58.8 | 53 | 54.5 | 85
| Std Dev. |  | 1.6432 | 0.5164 | 1.0954 | 0 | 1.5055 | 2.0412 | 3.3267 | 1.3416 | 0.73959 | 0.8165 | 1.3292 | 3.1464 | 1.169 | 1.0328 | 1.3292 | 1.0328 | 1.611 | 3.1885 | 1.0463 | 1.472 | 2.4685 | 0.98319 | 1.7889 | 0.75277
| Rel.Std.Dev. |  | 3.254% | 0.9332% | 2.067% | 0% | 2.771% | 3.45% | 5.871% | 2.396% | 1.248% | 1.87% | 2.409% | 5.569% | 2.094% | 1.936% | 2.5% | 1.801% | 4.043% | 5.389% | 1.922% | 1.749% | 4.129% | 1.849% | 3.252% | 0.8839%
| PDM3 |  | -8.67% | 0.07136% | -4.149% | -9.574% | -1.737% | 7.004% | 2.483% | 1.277% | 7.155% | -21.03% | -0.2301% | 2.181% | 0.9756% | -3.546% | -3.847% | 3.688% | -27.94% | 7.004% | -1.555% | 52.22% | 8.119% | -3.847% | -0.5315% | 54.03%
| Z-Score (Absolute) |  | 1.54 | 0.03 | 0.73 | 1.7 | 0.3 | 1.27 | 0.46 | 0.24 | 1.29 | 3.75 | 0.03 | 0.4 | 0.19 | 0.62 | 0.67 | 0.67 | 4.99 | 1.27 | 0.26 | 9.36 | 1.47 | 0.67 | 0.08 | 9.68
| Analytical results for Zr in OREAS 554 (Certified Value 63 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS |  |  | 0
1
| 1 | 1 | 59.5 | 63.5 | 65 | 44 | 61.8 | 60 | 61.3 | 66 | 60.2 | 56 | 64.5 | 64.8 | 65.6 | 65.4 | 63.8 | 65.7 | 77.7364 | 58.4 | 84.2566 | 63.8 | 58.8 | 61 | 77.9 | 63 |  |  | 1
| 1 | 2 | 58.4 | 65.6 | 62 | 39 | 61.8 | 60 | 60.6 | 68 | 58.8 | 57 | 65.7 | 65.5 | 66.9 | 66.9 | 67.2 | 65.1 | 77.5925 | 59.1 | 86.2228 | 63.1 | 60.5 | 61.6 | 75.9 | 64 |  |  | 0
| 1 | 3 | 59 | 63.4 | 62 | 38 | 61.8 | 60 | 62.6 | 71 | 60.5 | 56 | 64.3 | 62.6 | 64.7 | 65.1 | 64.4 | 64.9 | 81.3069 | 62.2 | 85.0947 | 61.3 | 58.6 | 61 | 74.6 | 65 |  |  | 16
| 1 | 4 | 59.4 | 65.1 | 62 | 39 | 62.3 | 60 | 60.8 | 68 | 59.9 | 55 | 62.9 | 64.1 | 65.2 | 67 | 66.3 | 64.2 | 77.6422 | 59 | 85.864 | 61.3 | 60.7 | 60.6 | 76.4 | 64 |  |  | 62.518
| 1 | 5 | 57.4 | 64.9 | 64 | 40 | 62.3 | 60 | 62.7 | 69 | 60.4 | 56 | 62.3 | 65.3 | 65.7 | 67 | 65.3 | 65.8 | 76.1423 | 60.2 | 82.8067 | 62.4 | 57.7 | 60.6 | 79.6 | 61 |  |  | 141
| 1 | 6 | 58.9 | 65.1 | 62 | 40 | 62.1 | 65 | 61.9 | 67 | 60.4 | 56 | 61.9 | 63.2 | 66.2 | 66.8 | 63.3 | 65.5 | 76.4012 | 59.6 | 86.0436 | 64.3 | 62 | 61.4 | 76.3 | 66
| Mean |  | 58.7667 | 64.6 | 62.8333 | 40 | 62.0167 | 60.8333 | 61.65 | 68.1667 | 60.0333 | 56 | 63.6 | 64.25 | 65.7167 | 66.3667 | 65.05 | 65.2 | 77.8036 | 59.75 | 85.0481 | 62.7 | 59.7167 | 61.0333 | 76.7833 | 63.8333
| Median |  | 58.95 | 65 | 62 | 39.5 | 61.95 | 60 | 61.6 | 68 | 60.3 | 56 | 63.6 | 64.45 | 65.65 | 66.85 | 64.85 | 65.3 | 77.6174 | 59.35 | 85.4793 | 62.75 | 59.65 | 61 | 76.35 | 64
| Std Dev. |  | 0.77632 | 0.92087 | 1.3292 | 2.0976 | 0.24833 | 2.0412 | 0.8961 | 1.7224 | 0.64083 | 0.63246 | 1.4683 | 1.1675 | 0.7679 | 0.87331 | 1.503 | 0.6 | 1.848 | 1.3442 | 1.3183 | 1.2602 | 1.6092 | 0.40825 | 1.7383 | 1.7224
| Rel.Std.Dev. |  | 1.321% | 1.425% | 2.115% | 5.244% | 0.4004% | 3.355% | 1.454% | 2.527% | 1.067% | 1.129% | 2.309% | 1.817% | 1.168% | 1.316% | 2.311% | 0.9202% | 2.375% | 2.25% | 1.55% | 2.01% | 2.695% | 0.6689% | 2.264% | 2.698%
| PDM3 |  | -6% | 3.33% | 0.5044% | -36.02% | -0.8019% | -2.695% | -1.388% | 9.035% | -3.974% | -10.43% | 1.731% | 2.77% | 5.116% | 6.156% | 4.05% | 4.29% | 24.45% | -4.428% | 36.04% | 0.2911% | -4.481% | -2.375% | 22.82% | 2.104%
| Z-Score (Absolute) |  | 1.23 | 0.38 | 0.11 | 6.41 | 0.33 | 0.66 | 0.43 | 1.37 | 0.88 | 1.99 | 0.11 | 0.29 | 0.69 | 0.87 | 0.51 | 0.55 | 4.03 | 0.96 | 6.03 | 0.14 | 0.97 | 0.6 | 3.75 | 0.17
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C330: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C348: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C367: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C385: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C478: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C496: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C533: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C552: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C570: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C607: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C625: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C644: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C680: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C698: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C716: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C734: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C752: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C770: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C788: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C807: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C826: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C844: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C862: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C918: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C937: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C956: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C974: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C992: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1011: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1030: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1048: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1066: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1084: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1103: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 554 (Certified Value 0.178 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.23 | 0.2 | 0.2 | 0.2 | 0.2 | < 0.5 | 0.175 | < 0.5 | 0.18 | 0.18 | 0.19 | 0.19 | 0.23 | 0.15 | 0.14762 | < 1 | 1.7 | < 0.2 | 0.19 | 0.155 | 0.19 | 0.17 | 0.18 |  |  | 1
| 1 | 2 | 0.25 | 0.18 | 0.2 | 0.21 | 0.2 | < 0.5 | 0.177 | < 0.5 | 0.17 | 0.19 | 0.19 | 0.19 | 0.18 | 0.13 | 0.16337 | < 1 | 1.64 | < 0.2 | 0.18 | 0.172 | 0.19 | 0.18 | 0.19 |  |  | 0
| 1 | 3 | 0.19 | 0.17 | 0.2 | 0.21 | 0.2 | < 0.5 | 0.17 | < 0.5 | 0.17 | 0.18 | 0.2 | 0.19 | 0.18 | 0.15 | 0.14134 | < 1 | 1.83 | < 0.2 | 0.19 | 0.155 | 0.18 | 0.17 | 0.19 |  |  | 16
| 1 | 4 | 0.18 | 0.18 | 0.2 | 0.2 | 0.2 | < 0.5 | 0.175 | < 0.5 | 0.17 | 0.19 | 0.19 | 0.19 | 0.17 | 0.13 | 0.13902 | < 1 | 1.7 | < 0.2 | 0.18 | 0.167 | 0.19 | 0.17 | 0.18 |  |  | 0.17772
| 1 | 5 | 0.19 | 0.17 | 0.2 | 0.2 | 0.2 | < 0.5 | 0.176 | < 0.5 | 0.17 | 0.17 | 0.18 | 0.17 | 0.18 | 0.13 | 0.16192 | < 1 | 1.81 | < 0.2 | 0.18 | 0.158 | 0.19 | 0.19 | 0.16 |  |  | 143
| 1 | 6 | 0.2 | 0.18 | 0.2 | 0.22 | 0.2 | < 0.5 | 0.163 | < 0.5 | 0.2 | 0.16 | 0.18 | 0.19 | 0.18 | 0.13 | 0.18774 | < 1 | 1.7 | < 0.2 | 0.18 | 0.175 | 0.18 | 0.17 | 0.2
| Mean |  | 0.20667 | 0.18 | 0.2 | 0.20667 | 0.2 |  | 0.17267 |  | 0.17667 | 0.17833 | 0.18833 | 0.18667 | 0.18667 | 0.13667 | 0.15683 |  | 1.73 |  | 0.18333 | 0.16367 | 0.18667 | 0.175 | 0.18333
| Median |  | 0.195 | 0.18 | 0.2 | 0.205 | 0.2 |  | 0.175 |  | 0.17 | 0.18 | 0.19 | 0.19 | 0.18 | 0.13 | 0.15477 |  | 1.7 |  | 0.18 | 0.1625 | 0.19 | 0.17 | 0.185
| Std Dev. |  | 0.027325 | 0.010954 | 3.0405e-17 | 0.008165 | 3.0405e-17 |  | 0.0053166 |  | 0.012111 | 0.01169 | 0.0075277 | 0.008165 | 0.021602 | 0.010328 | 0.018256 |  | 0.073756 |  | 0.005164 | 0.0088468 | 0.005164 | 0.0083666 | 0.013663
| Rel.Std.Dev. |  | 13.22% | 6.086% | 1.52e-14% | 3.951% | 1.52e-14% |  | 3.079% |  | 6.855% | 6.555% | 3.997% | 4.374% | 11.57% | 7.557% | 11.64% |  | 4.263% |  | 2.817% | 5.405% | 2.766% | 4.781% | 7.452%
| PDM3 |  | 16.29% | 1.284% | 12.54% | 16.29% | 12.54% |  | -2.842% |  | -0.5914% | 0.3464% | 5.973% | 5.035% | 5.035% | -23.1% | -11.75% |  | 873.5% |  | 3.16% | -7.906% | 5.035% | -1.529% | 3.16%
| Z-Score (Absolute) |  | 1.48 | 0.21 | NA | 1.48 | NA | 4.21 | 0.67 | 4.21 | 0.42 | 0.32 | 0.32 | 0.21 | 0.21 | 2.95 | 1.68 | 20.02 | 97.77 | 5.27 | 0 | 1.24 | 0.21 | 0.53 | 0
Results from laboratories 4 & 6 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 554 (Certified Value 0.594 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 27 | 28 | 29 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*AAS | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.6 | 0.625 | 0.61 | 0.684 | 0.5455 | 0.64 | 0.85 | 0.64 | 0.79 | 0.6 | 0.62 | 0.61 | 0.6 | 0.55 | NA | 0.57502 | 0.52 | 0.818 | 0.6 | 0.81571 | 0.65 | 0.55 |  |  |  | 1
| 1 | 2 | 0.62 | 0.6413 | 0.61 | 0.705 | 0.5195 | 0.62 | 0.82 | 0.63 | 0.8 | 0.61 | 0.6 | 0.61 | 0.62 | 0.55 | NA | 0.55211 | 0.53 | 0.827 | 0.64 | 0.83594 | 0.65 | 0.53 |  |  |  | 0
| 1 | 3 | 0.61 | 0.6082 | 0.58 | 0.792 | 0.5215 | 0.62 | 0.83 | 0.64 | 0.78 | 0.61 | 0.61 | 0.6 | 0.61 | 0.53 | NA | 0.54004 | 0.54 | 0.829 | 0.64 | 0.78629 | 0.63 | 0.52 |  |  |  | 16
| 1 | 4 | 0.58 | 0.5831 | 0.58 | 0.736 | 0.519 | 0.63 | 0.87 | 0.62 | 0.79 | 0.6 | 0.59 | 0.58 | 0.59 | 0.54 | NA | 0.56615 | 0.53 | 0.817 | 0.62 | 0.7466 | 0.65 | 0.52 |  |  |  | 0.5945
| 1 | 5 | 0.57 | 0.5725 | 0.58 | 0.779 | 0.525 | 0.63 | 0.82 | 0.62 | 0.79 | 0.59 | 0.57 | 0.58 | 0.57 | 0.56 | NA | 0.55566 | 0.55 | 0.816 | 0.58 | 0.8057 | 0.64 | 0.53 |  |  |  | 144
| 1 | 6 | 0.56 | 0.6249 | 0.58 | 0.744 | 0.5415 | 0.64 | 0.8 | 0.61 | 0.79 | 0.6 | 0.58 | 0.59 | 0.6 | 0.56 | NA | 0.57523 | 0.54 | 0.814 | 0.55 | 0.7561 | 0.6 | 0.54
| Mean |  | 0.59 | 0.60917 | 0.59 | 0.74 | 0.52867 | 0.63 | 0.83167 | 0.62667 | 0.79 | 0.60167 | 0.595 | 0.595 | 0.59833 | 0.54833 |  | 0.5607 | 0.535 | 0.82017 | 0.605 | 0.79105 | 0.63667 | 0.53167
| Median |  | 0.59 | 0.61655 | 0.58 | 0.74 | 0.52325 | 0.63 | 0.825 | 0.625 | 0.79 | 0.6 | 0.595 | 0.595 | 0.6 | 0.55 |  | 0.5609 | 0.535 | 0.8175 | 0.61 | 0.79599 | 0.645 | 0.53
| Std Dev. |  | 0.023664 | 0.026667 | 0.015492 | 0.041516 | 0.01175 | 0.0089443 | 0.024833 | 0.012111 | 0.0063246 | 0.0075277 | 0.018708 | 0.013784 | 0.017224 | 0.01169 |  | 0.01394 | 0.010488 | 0.0062423 | 0.035637 | 0.034806 | 0.019664 | 0.01169
| Rel.Std.Dev. |  | 4.011% | 4.378% | 2.626% | 5.61% | 2.223% | 1.42% | 2.986% | 1.933% | 0.8006% | 1.251% | 3.144% | 2.317% | 2.879% | 2.132% |  | 2.486% | 1.96% | 0.7611% | 5.89% | 4.4% | 3.089% | 2.199%
| PDM3 |  | -0.7565% | 2.468% | -0.7565% | 24.47% | -11.07% | 5.972% | 39.89% | 5.411% | 32.89% | 1.206% | 0.08459% | 0.08459% | 0.6453% | -7.765% |  | -5.685% | -10.01% | 37.96% | 1.767% | 33.06% | 7.093% | -10.57%
| Z-Score (Absolute) |  | 0.22 | 0.14 | 0.22 | 2.67 | 1.41 | 0.55 | 4.43 | 0.48 | 3.63 | 0 | 0.13 | 0.13 | 0.06 | 1.03 | NA | 0.79 | 1.28 | 4.21 | 0.06 | 3.65 | 0.67 | 1.35
| Analytical results for As in OREAS 554 (Certified Value 65 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  | 0
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 1
| 1 | 1 | 63 | 62 | 53 | 60 | 61 | 68 | 62 | 67.4 | 66 | 70.6 | 71.6 | 65.3 | 70.3 | 70.6 | 51.9 | 55.4207 | 65 | 66.1533 | 65 | 62 | 62.7 | 66 | 47.6 |  |  | 1
| 1 | 2 | 64 | 63 | 54 | 62 | 62 | 70 | 61.2 | 66.1 | 66 | 70.3 | 69.1 | 64.5 | 68.1 | 70 | 52.5 | 62.0491 | 64 | 66.041 | 64 | 63.2 | 64.6 | 67 | 45.9 |  |  | 0
| 1 | 3 | 64 | 62 | 53 | 60 | 66 | 67 | 61 | 68.4 | 64 | 71.5 | 70.7 | 64.2 | 67.7 | 69.9 | 52.8 | 53.4652 | 66 | 65.7193 | 63 | 63.2 | 61.6 | 64 | 50.9 |  |  | 16
| 1 | 4 | 61 | 61 | 53.5 | 61 | 65 | 69 | 62.3 | 68.7 | 66 | 69.6 | 70.7 | 63.7 | 66.9 | 64.7 | 52.4 | 57.0298 | 66 | 65.349 | 64 | 62.9 | 63.3 | 65 | 48.6 |  |  | 64.8996
| 1 | 5 | 62 | 60 | 53.5 | 63 | 66 | 69 | 63.6 | 67.5 | 65 | 67.8 | 68.3 | 64.6 | 62.1 | 64.9 | 53.2 | 58.0674 | 66 | 66.5213 | 63 | 60.9 | 62.7 | 66 | 48.5 |  |  | 145
| 1 | 6 | 63 | 62 | 53 | 63 | 64 | 69 | 62.4 | 68.6 | 65 | 69.4 | 69.5 | 66.1 | 68.7 | 66 | 53.1 | 57.9521 | 65 | 66.5082 | 64 | 60.3 | 60.8 | 65 | 47.8
| Mean |  | 62.8333 | 61.6667 | 53.3333 | 61.5 | 64 | 68.6667 | 62.0833 | 67.7833 | 65.3333 | 69.8667 | 69.9833 | 64.7333 | 67.3 | 67.6833 | 52.65 | 57.3307 | 65.3333 | 66.0487 | 63.8333 | 62.0833 | 62.6167 | 65.5 | 48.2167
| Median |  | 63 | 62 | 53.25 | 61.5 | 64.5 | 69 | 62.15 | 67.95 | 65.5 | 69.95 | 70.1 | 64.55 | 67.9 | 67.95 | 52.65 | 57.491 | 65.5 | 66.0972 | 64 | 62.45 | 62.7 | 65.5 | 48.15
| Std Dev. |  | 1.169 | 1.0328 | 0.40825 | 1.3784 | 2.0976 | 1.0328 | 0.93897 | 0.99482 | 0.8165 | 1.2612 | 1.2238 | 0.8501 | 2.7914 | 2.7665 | 0.48477 | 2.8939 | 0.8165 | 0.45685 | 0.75277 | 1.2449 | 1.3227 | 1.0488 | 1.6339
| Rel.Std.Dev. |  | 1.861% | 1.675% | 0.7655% | 2.241% | 3.278% | 1.504% | 1.512% | 1.468% | 1.25% | 1.805% | 1.749% | 1.313% | 4.148% | 4.087% | 0.9207% | 5.048% | 1.25% | 0.6917% | 1.179% | 2.005% | 2.112% | 1.601% | 3.389%
| PDM3 |  | -3.184% | -4.981% | -17.82% | -5.238% | -1.386% | 5.804% | -4.339% | 4.443% | 0.6683% | 7.653% | 7.833% | -0.2562% | 3.699% | 4.289% | -18.87% | -11.66% | 0.6683% | 1.771% | -1.643% | -4.339% | -3.518% | 0.9251% | -25.71%
| Z-Score (Absolute) |  | 0.34 | 0.67 | 3.08 | 0.72 | 0 | 1.35 | 0.55 | 1.09 | 0.39 | 1.7 | 1.73 | 0.21 | 0.95 | 1.06 | 3.28 | 1.93 | 0.39 | 0.59 | 0.05 | 0.55 | 0.4 | 0.43 | 4.56
| Analytical results for Au in OREAS 554 (Indicative Value 0.013 ppm) | [26 empty] | Indicative
Au | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab |  |  | 1
| No | No | 01: 02; 02: 04; 04: 05; 05: 06; 06: 10; 09: 13; 10: 14; 11: 15; 12: 16; 13: 17; 14: 22; 15: 24; 16: 28 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 09: AR*MS; 10: AR*MS; 11: AR*MS; 12: AR*MS; 13: AR*MS; 14: AR*MS; 15: AR*MS; 16: AR*OES/MS |  |  | 3
|  |  | 01: 25g; 02: 01-05g; 04: 01g; 05: 25g; 06: 0.5g; 09: 0.5g; 10: 0.5g; 11: 0.5g; 12: 0.5g; 13: 0.5g; 14: 0.5g; 15: 0.5g; 16: 0.5g |  |  | 3
| 1 | 1 | 01: 0.028; 02: 0.019; 04: < 0.001; 05: 0.01; 06: 0.0006; 09: < 0.02; 10: < 0.02; 11: < 0.02; 12: < 0.02; 13: < 0.02; 14: 0.02581; 15: 0.0073; 16: 0.0064 |  |  | 1
| 1 | 2 | 01: 0.027; 02: 0.018; 04: < 0.001; 05: 0.01; 06: 0.0011; 09: < 0.02; 10: < 0.02; 11: < 0.02; 12: < 0.02; 13: < 0.02; 14: 0.02268; 15: 0.0083; 16: 0.006 |  |  | 0
| 1 | 3 | 01: 0.028; 02: 0.016; 04: < 0.001; 05: 0.01; 06: 0.0017; 09: < 0.02; 10: < 0.02; 11: < 0.02; 12: < 0.02; 13: < 0.02; 14: 0.01819; 15: 0.0068; 16: 0.0057 |  |  | 16
| 1 | 4 | 01: 0.027; 02: 0.017; 04: < 0.001; 05: 0.01; 06: 0.001; 09: < 0.02; 10: < 0.02; 11: < 0.02; 12: < 0.02; 13: < 0.02; 14: 0.02158; 15: 0.008; 16: 0.0063 |  |  | 0.01307
| 1 | 5 | 01: 0.028; 02: 0.017; 04: < 0.001; 05: 0.01; 06: < 0.0005; 09: < 0.02; 10: < 0.02; 11: < 0.02; 12: < 0.02; 13: < 0.02; 14: 0.02011; 15: 0.008; 16: 0.0063 |  |  | 27
| 1 | 6 | 01: 0.026; 02: 0.019; 04: < 0.001; 05: 0.01; 06: 0.001; 09: < 0.02; 10: < 0.02; 11: < 0.02; 12: < 0.02; 13: < 0.02; 14: 0.02031; 15: 0.0082; 16: 0.0065
| Mean |  | 01: 0.027333; 02: 0.017667; 05: 0.01; 06: 0.00108; 14: 0.021447; 15: 0.0077667; 16: 0.0062
| Median |  | 01: 0.0275; 02: 0.0175; 05: 0.01; 06: 0.001; 14: 0.020945; 15: 0.008; 16: 0.0063
| Std Dev. |  | 01: 0.0008165; 02: 0.0012111; 05: 0; 06: 0.00039623; 14: 0.0026177; 15: 0.00058878; 16: 0.00029665
| Rel.Std.Dev. |  | 01: 2.987%; 02: 6.855%; 05: 0%; 06: 36.69%; 14: 12.21%; 15: 7.581%; 16: 4.785%
| PDM3 |  | 01: 109.1%; 02: 35.16%; 05: -23.49%; 06: -91.74%; 14: 64.08%; 15: -40.58%; 16: -52.56%
| Z-Score (Absolute) |  | 01: 5.23; 02: 2.31; 04: 2.87; 05: 0; 06: 2.73; 09: 0; 10: 0; 11: 0; 12: 0; 13: 0; 14: 3.46; 15: 0.67; 16: 1.15
| Analytical results for B in OREAS 554 (Certified Value < 10 ppm) | [26 empty] | Certified
B | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab |  |  | 1
| No | No | 01: 02; 02: 05; 04: 10; 05: 12; 06: 13; 09: 14; 10: 15; 11: 16; 12: 17; 13: 21; 14: 24; 15: 28; 16: 29 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*OES/MS; 05: AR*OES; 06: AR*MS; 09: AR*MS; 10: AR*MS; 11: AR*MS; 12: AR*MS; 13: AR*OES/MS; 14: AR*MS; 15: AR*OES/MS; 16: AR*OES/MS |  |  | 1
|  |  | 01: 25g; 02: 01g; 04: 0.5g; 05: 0.5g; 06: 0.5g; 09: 0.5g; 10: 0.5g; 11: 0.5g; 12: 0.5g; 13: 0.25g; 14: 0.5g; 15: 0.5g; 16: 0.25g |  |  | 1
| 1 | 1 | 01: < 10; 02: < 10; 04: 7; 05: 10; 06: < 10; 09: < 10; 10: < 10; 11: < 10; 12: < 10; 13: < 10; 14: < 20; 15: < 10; 16: < 10 |  |  | 1
| 1 | 2 | 01: < 10; 02: < 10; 04: 7; 05: 11; 06: < 10; 09: 10; 10: < 10; 11: < 10; 12: < 10; 13: < 10; 14: < 20; 15: < 10; 16: < 10 |  |  | 0
| 1 | 3 | 01: < 10; 02: < 10; 04: 8; 05: 13; 06: < 10; 09: 10; 10: < 10; 11: < 10; 12: < 10; 13: < 10; 14: < 20; 15: < 10; 16: < 10 |  |  | 16
| 1 | 4 | 01: < 10; 02: < 10; 04: 7; 05: 12; 06: < 10; 09: < 10; 10: < 10; 11: < 10; 12: < 10; 13: < 10; 14: < 20; 15: < 10; 16: < 10 |  |  | < 10
| 1 | 5 | 01: < 10; 02: < 10; 04: 8; 05: 12; 06: < 10; 09: < 10; 10: < 10; 11: < 10; 12: < 10; 13: < 10; 14: < 20; 15: < 10; 16: < 10 |  |  | 146
| 1 | 6 | 01: < 10; 02: < 10; 04: 7; 05: 12; 06: < 10; 09: < 10; 10: < 10; 11: < 10; 12: < 10; 13: < 10; 14: < 20; 15: < 10; 16: < 10
| Mean |  | 04: 7.3333; 05: 11.6667; 09: 10
| Median |  | 04: 7; 05: 12; 09: 10
| Std Dev. |  | 04: 0.5164; 05: 1.0328; 09: 0
| Rel.Std.Dev. |  | 04: 7.042%; 05: 8.853%; 09: 0%
PDM3
Z-Score (Absolute) | bar: 01 to 16
| Analytical results for Ba in OREAS 554 (Certified Value 226 ppm) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | [5 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [5 empty] | 0
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [5 empty] | 0
| 1 | 1 | 209 | 222 | 199 | 246 | 222 | 273 | 39.1 | 144 | 220 | 230 | 230 | 230 | 220 | NA | 263.1206 | 196 | 287 | 233.1 | 233 | 217 | [5 empty] | 1
| 1 | 2 | 219 | 210 | 213 | 249 | 213 | 266 | 43 | 144 | 220 | 230 | 230 | 240 | 210 | NA | 274.2604 | 202 | 286 | 242.4 | 238 | 215 | [5 empty] | 0
| 1 | 3 | 215 | 185 | 192 | 269 | 215 | 270 | 32.5 | 144 | 230 | 230 | 220 | 230 | 210 | NA | 251.8373 | 206 | 291 | 244.3 | 232 | 215 | [5 empty] | 16
| 1 | 4 | 201 | 208 | 188 | 263 | 214 | 274 | 30.2 | 147 | 220 | 220 | 210 | 230 | 210 | NA | 259.3277 | 197 | 281 | 241.4 | 237 | 213 | [5 empty] | 225.9637
| 1 | 5 | 193 | 217 | 188 | 270 | 218 | 278 | 33.7 | 145 | 220 | 210 | 210 | 220 | 210 | NA | 250.418 | 206 | 278 | 227.1 | 233 | 219 | [5 empty] | 147
| 1 | 6 | 190 | 219 | 189 | 261 | 221 | 273 | 35.2 | 143 | 220 | 220 | 220 | 230 | 210 | NA | 273.0498 | 206 | 278 | 218.2 | 219 | 214
| Mean |  | 204.5 | 210.1667 | 194.8333 | 259.6667 | 217.1667 | 272.3333 | 35.6167 | 144.5 | 221.6667 | 223.3333 | 220 | 230 | 211.6667 |  | 262.0023 | 202.1667 | 283.5 | 234.4167 | 232 | 215.5
| Median |  | 205 | 213.5 | 190.5 | 262 | 216.5 | 273 | 34.45 | 144 | 220 | 225 | 220 | 230 | 210 |  | 261.2241 | 204 | 283.5 | 237.25 | 233 | 215
| Std Dev. |  | 11.7941 | 13.4375 | 9.8268 | 10.0731 | 3.7639 | 4.0332 | 4.6808 | 1.3784 | 4.0825 | 8.165 | 8.9443 | 6.3246 | 4.0825 |  | 10.1807 | 4.6655 | 5.3198 | 10.2804 | 6.8118 | 2.1679
| Rel.Std.Dev. |  | 5.767% | 6.394% | 5.044% | 3.879% | 1.733% | 1.481% | 13.14% | 0.9539% | 1.842% | 3.656% | 4.066% | 2.75% | 1.929% |  | 3.886% | 2.308% | 1.876% | 4.386% | 2.936% | 1.006%
| PDM3 |  | -9.499% | -6.991% | -13.78% | 14.92% | -3.893% | 20.52% | -84.24% | -36.05% | -1.902% | -1.164% | -2.639% | 1.786% | -6.327% |  | 15.95% | -10.53% | 25.46% | 3.741% | 2.671% | -4.631%
| Z-Score (Absolute) |  | 0.72 | 0.46 | 1.18 | 1.86 | 0.13 | 2.45 | 8.62 | 3.53 | 0.08 | 0.16 | 0 | 0.47 | 0.39 | NA | 1.96 | 0.83 | 2.97 | 0.67 | 0.56 | 0.21
| Analytical results for Be in OREAS 554 (Certified Value 0.63 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | 30 | [5 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.7 | 0.67 | < 1 | 0.7 | 0.7 | 0.6 | 0.6 | < 5 | 0.53 | 0.64 | 0.63 | 0.61 | 0.59 | 0.37 | 1.2121 | 0.6 | 0.68 | 0.63 | 0.6 | 0.64 | [5 empty] | 1
| 1 | 2 | 0.7 | 0.65 | < 1 | 0.6 | 0.6 | 0.7 | 0.6 | < 5 | 0.54 | 0.67 | 0.63 | 0.62 | 0.62 | 0.37 | 1.3904 | 0.6 | 0.69 | 0.67 | 0.5 | 0.59 | [5 empty] | 0
| 1 | 3 | 0.7 | 0.62 | < 1 | 0.7 | 0.7 | 0.6 | 0.7 | < 5 | 0.53 | 0.65 | 0.56 | 0.61 | 0.63 | 0.37 | 1.1306 | 0.6 | 0.69 | 0.6 | 0.6 | 0.52 | [5 empty] | 16
| 1 | 4 | 0.6 | 0.61 | < 1 | 0.6 | 0.6 | 0.6 | 0.6 | < 5 | 0.53 | 0.69 | 0.62 | 0.61 | 0.64 | 0.37 | 1.2142 | 0.6 | 0.69 | 0.67 | 0.7 | 0.64 | [5 empty] | 0.62889
| 1 | 5 | 0.6 | 0.63 | < 1 | 0.7 | 0.6 | 0.6 | 0.6 | < 5 | 0.5 | 0.63 | 0.62 | 0.58 | 0.62 | 0.37 | 1.2821 | 0.6 | 0.66 | 0.63 | 0.6 | 0.52 | [5 empty] | 148
| 1 | 6 | 0.6 | 0.62 | < 1 | 0.7 | 0.6 | 0.6 | 0.6 | < 5 | 0.52 | 0.67 | 0.63 | 0.62 | 0.6 | 0.37 | 1.3386 | 0.6 | 0.68 | 0.64 | 0.6 | 0.57
| Mean |  | 0.65 | 0.63333 |  | 0.66667 | 0.63333 | 0.61667 | 0.61667 |  | 0.525 | 0.65833 | 0.615 | 0.60833 | 0.61667 | 0.37 | 1.2613 | 0.6 | 0.68167 | 0.64 | 0.6 | 0.58
| Median |  | 0.65 | 0.625 |  | 0.7 | 0.6 | 0.6 | 0.6 |  | 0.53 | 0.66 | 0.625 | 0.61 | 0.62 | 0.37 | 1.2482 | 0.6 | 0.685 | 0.635 | 0.6 | 0.58
| Std Dev. |  | 0.054772 | 0.022509 |  | 0.05164 | 0.05164 | 0.040825 | 0.040825 |  | 0.013784 | 0.022286 | 0.027386 | 0.01472 | 0.018619 | 6.0809e-17 | 0.09467 | 0 | 0.01169 | 0.026833 | 0.063246 | 0.054037
| Rel.Std.Dev. |  | 8.427% | 3.554% |  | 7.746% | 8.154% | 6.62% | 6.62% |  | 2.626% | 3.385% | 4.453% | 2.42% | 3.019% | 1.643e-14% | 7.505% | 0% | 1.715% | 4.193% | 10.54% | 9.317%
| PDM3 |  | 3.357% | 0.7067% |  | 6.007% | 0.7067% | -1.943% | -1.943% |  | -16.52% | 4.682% | -2.208% | -3.269% | -1.943% | -41.17% | 100.6% | -4.594% | 8.392% | 1.767% | -4.594% | -7.774%
| Z-Score (Absolute) |  | 0.79 | 0.4 | 2.78 | 1.19 | 0.4 | 0 | 0 | 44.82 | 2.18 | 0.99 | 0.04 | 0.2 | 0 | 5.87 | 15.34 | 0.4 | 1.55 | 0.56 | 0.4 | 0.87
| Analytical results for Bi in OREAS 554 (Certified Value 2.58 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 2.42 | 2.64 | 2.5 | 2.34 | 2.7 | 2.9 | 2.59 | 3.38 | 2 | 2.72 | 2.56 | 2.28 | 2.69 | 2.54 | 3.17 | 2.2947 | 2.51 | 2.5639 | 2.57 | 2.43 | 2.5 | 2.54 | 2.65 |  |  | 1
| 1 | 2 | 2.41 | 2.72 | 2.6 | 2.41 | 2.6 | 2.9 | 2.66 | 3.09 | 2 | 2.64 | 2.67 | 2.21 | 2.75 | 2.51 | 3.31 | 3.279 | 2.56 | 2.5611 | 2.68 | 2.53 | 2.57 | 2.56 | 2.9 |  |  | 0
| 1 | 3 | 2.36 | 2.66 | 2.5 | 2.45 | 2.6 | 2.9 | 2.53 | 3.23 | 2 | 2.67 | 2.61 | 2.11 | 2.76 | 2.49 | 3.53 | 2.6793 | 2.61 | 2.6286 | 2.62 | 2.62 | 2.46 | 2.53 | 2.78 |  |  | 16
| 1 | 4 | 2.35 | 2.62 | 2.4 | 2.33 | 2.8 | 2.9 | 2.64 | 3.15 | 2 | 2.56 | 2.6 | 2.28 | 2.72 | 2.52 | 3.33 | 2.3882 | 2.54 | 2.5752 | 2.54 | 2.4 | 2.54 | 2.52 | 2.8 |  |  | 2.5797
| 1 | 5 | 2.38 | 2.63 | 2.6 | 2.41 | 2.7 | 3 | 2.67 | 3.1 | 2 | 2.65 | 2.54 | 2.15 | 2.73 | 2.69 | 3.23 | 2.6688 | 2.54 | 2.5857 | 2.58 | 2.66 | 2.5 | 2.52 | 2.67 |  |  | 149
| 1 | 6 | 2.5 | 2.69 | 2.5 | 2.45 | 2.8 | 2.8 | 2.64 | 3.66 | 2 | 2.62 | 2.58 | 2.16 | 2.78 | 2.52 | 3.21 | 2.6831 | 2.54 | 2.5974 | 2.62 | 2.63 | 2.43 | 2.57 | 2.76
| Mean |  | 2.4033 | 2.66 | 2.5167 | 2.3983 | 2.7 | 2.9 | 2.6217 | 3.2683 | 2 | 2.6433 | 2.5933 | 2.1983 | 2.7383 | 2.545 | 3.2967 | 2.6655 | 2.55 | 2.5853 | 2.6017 | 2.545 | 2.5 | 2.54 | 2.76
| Median |  | 2.395 | 2.65 | 2.5 | 2.41 | 2.7 | 2.9 | 2.64 | 3.19 | 2 | 2.645 | 2.59 | 2.185 | 2.74 | 2.52 | 3.27 | 2.6741 | 2.54 | 2.5805 | 2.6 | 2.575 | 2.5 | 2.535 | 2.77
| Std Dev. |  | 0.05465 | 0.038471 | 0.075277 | 0.052313 | 0.089443 | 0.063246 | 0.052694 | 0.21977 | 0 | 0.053166 | 0.045461 | 0.070828 | 0.031885 | 0.07287 | 0.12941 | 0.34389 | 0.033466 | 0.025191 | 0.04916 | 0.11005 | 0.05099 | 0.020976 | 0.091433
| Rel.Std.Dev. |  | 2.274% | 1.446% | 2.991% | 2.181% | 3.313% | 2.181% | 2.01% | 6.724% | 0% | 2.011% | 1.753% | 3.222% | 1.164% | 2.863% | 3.925% | 12.9% | 1.312% | 0.9744% | 1.89% | 4.324% | 2.04% | 0.8258% | 3.313%
| PDM3 |  | -6.836% | 3.114% | -2.442% | -7.03% | 4.664% | 12.42% | 1.628% | 26.7% | -22.47% | 2.468% | 0.5295% | -14.78% | 6.15% | -1.344% | 27.79% | 3.327% | -1.15% | 0.2192% | 0.8525% | -1.344% | -3.089% | -1.538% | 6.99%
| Z-Score (Absolute) |  | 1.76 | 0.57 | 0.73 | 1.8 | 0.93 | 2.74 | 0.22 | 6.08 | NA | 0.42 | 0.04 | 3.62 | 1.28 | 0.48 | 6.33 | 0.62 | 0.43 | 0.11 | 0.04 | 0.48 | 0.88 | 0.52 | 1.47
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 554 (Certified Value 3.85 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 27 | 28 | 29 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*AAS | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 3.83 | 3.87 | 3.87 | 3.58 | 4.029 | 3.96 | 3.6 | 3.98 | 3.95 | 3.91 | 4.01 | 4.03 | 4.01 | 3.87 | 3.8129 | 4.146 | 3.95 | 3.75 | 3.69 | 3.7305 | 3.75 | 3.9 |  |  |  | 1
| 1 | 2 | 3.88 | 3.93 | 3.83 | 3.61 | 3.985 | 3.88 | 3.61 | 3.89 | 3.95 | 3.92 | 3.95 | 4.05 | 3.95 | 3.85 | 3.7883 | 4.035 | 3.93 | 3.74 | 3.76 | 3.813 | 3.83 | 3.92 |  |  |  | 0
| 1 | 3 | 3.84 | 3.89 | 3.82 | 3.67 | 4.02 | 3.95 | 3.53 | 4.07 | 3.99 | 4 | 3.98 | 3.98 | 3.98 | 3.81 | 3.9795 | 3.995 | 3.98 | 3.73 | 3.79 | 3.7723 | 3.75 | 3.83 |  |  |  | 16
| 1 | 4 | 3.68 | 3.75 | 3.93 | 3.67 | 4.002 | 3.96 | 3.56 | 3.91 | 3.99 | 3.92 | 3.83 | 3.95 | 3.87 | 3.74 | 3.8084 | 3.943 | 3.97 | 3.72 | 3.71 | 3.7612 | 3.79 | 3.85 |  |  |  | 3.8542
| 1 | 5 | 3.77 | 3.75 | 3.88 | 3.69 | 3.9745 | 3.97 | 3.73 | 3.91 | 4.01 | 3.83 | 3.81 | 3.93 | 3.84 | 3.77 | 3.8377 | 3.952 | 4.03 | 3.75 | 3.67 | 3.782 | 3.78 | 3.88 |  |  |  | 150
| 1 | 6 | 3.81 | 3.86 | 3.88 | 3.61 | 3.9965 | 3.96 | 3.6 | 3.9 | 4.07 | 3.95 | 3.86 | 3.97 | 4.04 | 3.84 | 3.8459 | 3.955 | 4.02 | 3.76 | 3.63 | 3.7213 | 3.68 | 3.88
| Mean |  | 3.8017 | 3.8417 | 3.8683 | 3.6383 | 4.0012 | 3.9467 | 3.605 | 3.9433 | 3.9933 | 3.9217 | 3.9067 | 3.985 | 3.9483 | 3.8133 | 3.8455 | 4.0043 | 3.98 | 3.7417 | 3.7083 | 3.7634 | 3.7633 | 3.8767
| Median |  | 3.82 | 3.865 | 3.875 | 3.64 | 3.9992 | 3.96 | 3.6 | 3.91 | 3.99 | 3.92 | 3.905 | 3.975 | 3.965 | 3.825 | 3.8253 | 3.975 | 3.975 | 3.745 | 3.7 | 3.7668 | 3.765 | 3.88
| Std Dev. |  | 0.06969 | 0.074944 | 0.039707 | 0.044008 | 0.020612 | 0.033267 | 0.068337 | 0.069761 | 0.044572 | 0.055648 | 0.084063 | 0.046368 | 0.078846 | 0.050067 | 0.068888 | 0.0775 | 0.038987 | 0.01472 | 0.058793 | 0.033909 | 0.050465 | 0.03266
| Rel.Std.Dev. |  | 1.833% | 1.951% | 1.026% | 1.21% | 0.5152% | 0.8429% | 1.896% | 1.769% | 1.116% | 1.419% | 2.152% | 1.164% | 1.997% | 1.313% | 1.791% | 1.935% | 0.9796% | 0.3934% | 1.585% | 0.901% | 1.341% | 0.8425%
| PDM3 |  | -1.363% | -0.325% | 0.3669% | -5.601% | 3.813% | 2.399% | -6.465% | 2.313% | 3.61% | 1.751% | 1.362% | 3.394% | 2.443% | -1.06% | -0.2268% | 3.896% | 3.264% | -2.92% | -3.784% | -2.356% | -2.357% | 0.5831%
| Z-Score (Absolute) |  | 0.52 | 0.23 | 0.03 | 1.72 | 0.95 | 0.55 | 1.97 | 0.52 | 0.89 | 0.36 | 0.25 | 0.83 | 0.56 | 0.44 | 0.2 | 0.97 | 0.79 | 0.96 | 1.21 | 0.8 | 0.8 | 0.03
| Analytical results for Cd in OREAS 554 (Certified Value 0.2 ppm) | [26 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.23 | 0.19 | 0.2 | 0.18 | < 0.5 | < 0.5 | 0.22 | 0.24 | < 0.5 | 0.2 | 0.21 | 0.21 | 0.19 | 0.2 | 0.14 | 0.2202 | 0.22 | 0.20886 | 0.2 | 0.18 | 0.21 | 0.21 | 0.19 |  |  | 1
| 1 | 2 | 0.25 | 0.18 | 0.2 | 0.17 | < 0.5 | < 0.5 | 0.18 | 0.21 | < 0.5 | 0.21 | 0.21 | 0.22 | 0.2 | 0.21 | 0.14 | 0.19623 | 0.2 | 0.2176 | 0.19 | 0.19 | 0.21 | 0.2 | 0.18 |  |  | 0
| 1 | 3 | 0.22 | 0.19 | 0.2 | 0.18 | < 0.5 | < 0.5 | 0.2 | 0.26 | < 0.5 | 0.21 | 0.22 | 0.21 | 0.2 | 0.2 | 0.16 | 0.17127 | 0.18 | 0.23371 | 0.19 | 0.16 | 0.19 | 0.2 | 0.19 |  |  | 16
| 1 | 4 | 0.19 | 0.17 | 0.2 | 0.17 | < 0.5 | < 0.5 | 0.22 | 0.2 | < 0.5 | 0.19 | 0.2 | 0.21 | 0.2 | 0.2 | 0.16 | 0.17256 | 0.19 | 0.21192 | 0.19 | 0.19 | 0.2 | 0.22 | 0.22 |  |  | 0.19873
| 1 | 5 | 0.21 | 0.2 | 0.2 | 0.17 | < 0.5 | < 0.5 | 0.17 | 0.18 | < 0.5 | 0.2 | 0.21 | 0.21 | 0.2 | 0.2 | 0.14 | 0.18374 | 0.21 | 0.20675 | 0.19 | 0.19 | 0.2 | 0.19 | 0.19 |  |  | 151
| 1 | 6 | 0.21 | 0.22 | 0.2 | 0.18 | < 0.5 | < 0.5 | 0.15 | 0.17 | < 0.5 | 0.19 | 0.19 | 0.2 | 0.22 | 0.22 | 0.14 | 0.19475 | 0.2 | 0.14954 | 0.19 | 0.21 | 0.19 | 0.17 | 0.21
| Mean |  | 0.21833 | 0.19167 | 0.2 | 0.175 |  |  | 0.19 | 0.21 |  | 0.2 | 0.20667 | 0.21 | 0.20167 | 0.205 | 0.14667 | 0.18979 | 0.2 | 0.20473 | 0.19167 | 0.18667 | 0.2 | 0.19833 | 0.19667
| Median |  | 0.215 | 0.19 | 0.2 | 0.175 |  |  | 0.19 | 0.205 |  | 0.2 | 0.21 | 0.21 | 0.2 | 0.2 | 0.14 | 0.18925 | 0.2 | 0.21039 | 0.19 | 0.19 | 0.2 | 0.2 | 0.19
| Std Dev. |  | 0.020412 | 0.017224 | 3.0405e-17 | 0.0054772 |  |  | 0.028284 | 0.034641 |  | 0.0089443 | 0.010328 | 0.0063246 | 0.0098319 | 0.0083666 | 0.010328 | 0.018261 | 0.014142 | 0.02872 | 0.0040825 | 0.01633 | 0.0089443 | 0.017224 | 0.015055
| Rel.Std.Dev. |  | 9.349% | 8.986% | 1.52e-14% | 3.13% |  |  | 14.89% | 16.5% |  | 4.472% | 4.997% | 3.012% | 4.875% | 4.081% | 7.042% | 9.621% | 7.071% | 14.03% | 2.13% | 8.748% | 4.472% | 8.684% | 7.655%
| PDM3 |  | 9.862% | -3.557% | 0.6366% | -11.94% |  |  | -4.395% | 5.668% |  | 0.6366% | 3.991% | 5.668% | 1.475% | 3.153% | -26.2% | -4.5% | 0.6366% | 3.017% | -3.557% | -6.072% | 0.6366% | -0.202% | -1.041%
| Z-Score (Absolute) |  | 1.35 | 0.61 | NA | 1.84 | 3.68 | 3.68 | 0.74 | 0.74 | 3.68 | 0 | 0.49 | 0.74 | 0.12 | 0.37 | 3.92 | 0.75 | 0 | 0.35 | 0.61 | 0.98 | 0 | 0.12 | 0.25
Results from laboratory 4 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 554 (Certified Value 34.8 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | [7 empty] | 1
| 1 | 1 | 33.55 | 32.217 | 27.5 | 35.78 | 33.5 | 38.4 | 35.8 | 36.4 | 36.3 | 37.5 | 36.1 | 33.9 | 24.2 | 34.8427 | 29.26 | 42 | 40.48 | 32.39 | [7 empty] | 1
| 1 | 2 | 33.24 | 33.333 | 26.7 | 35.57 | 31.5 | 39.3 | 34.8 | 37.2 | 35.4 | 37.8 | 37.1 | 32.7 | 24.5 | 33.9094 | 28.87 | 43.4 | 40.86 | 32.83 | [7 empty] | 0
| 1 | 3 | 34.23 | 32.294 | 25.6 | 36.88 | 31 | 38.1 | 35.2 | 37.4 | 35.7 | 35.6 | 37.5 | 32.8 | 27.3 | 33.0803 | 30.23 | 42.4 | 39.28 | 32.48 | [7 empty] | 16
| 1 | 4 | 32.74 | 31.158 | 26.1 | 36.82 | 31.5 | 38.5 | 33.6 | 36.2 | 35.6 | 34.7 | 37.1 | 32.9 | 26 | 35.5744 | 30.72 | 41.6 | 40.35 | 32.31 | [7 empty] | 34.7548
| 1 | 5 | 32.4 | 31.598 | 25.6 | 37.43 | 30 | 38.5 | 34.3 | 36 | 34.2 | 35.5 | 39.3 | 32.9 | 25.3 | 33.4606 | 30.54 | 42.8 | 39.85 | 33.11 | [7 empty] | 152
| 1 | 6 | 31.95 | 32.396 | 25.9 | 37.52 | 33 | 39.2 | 34.7 | 36.8 | 34.8 | 35.9 | 37.5 | 32.9 | 24.7 | 36.0321 | 30.39 | 42 | 38 | 31.89
| Mean |  | 33.0183 | 32.166 | 26.2333 | 36.6667 | 31.75 | 38.6667 | 34.7333 | 36.6667 | 35.3333 | 36.1667 | 37.4333 | 33.0167 | 25.3333 | 34.4832 | 30.0017 | 42.3667 | 39.8033 | 32.5017
| Median |  | 32.99 | 32.2555 | 26 | 36.85 | 31.5 | 38.5 | 34.75 | 36.6 | 35.5 | 35.75 | 37.3 | 32.9 | 25 | 34.376 | 30.31 | 42.2 | 40.1 | 32.435
| Std Dev. |  | 0.82475 | 0.74464 | 0.74207 | 0.82082 | 1.2942 | 0.4761 | 0.75277 | 0.56095 | 0.73666 | 1.2193 | 1.0482 | 0.44008 | 1.157 | 1.1884 | 0.75359 | 0.65013 | 1.0394 | 0.42466
| Rel.Std.Dev. |  | 2.498% | 2.315% | 2.829% | 2.239% | 4.076% | 1.231% | 2.167% | 1.53% | 2.085% | 3.371% | 2.8% | 1.333% | 4.567% | 3.446% | 2.512% | 1.535% | 2.611% | 1.307%
| PDM3 |  | -4.996% | -7.449% | -24.52% | 5.501% | -8.646% | 11.26% | -0.06165% | 5.501% | 1.665% | 4.062% | 7.707% | -5.001% | -27.11% | -0.7813% | -13.68% | 21.9% | 14.53% | -6.483%
| Z-Score (Absolute) |  | 0.47 | 0.72 | 2.48 | 0.61 | 0.85 | 1.2 | 0.04 | 0.61 | 0.21 | 0.46 | 0.84 | 0.47 | 2.75 | 0.04 | 1.37 | 2.3 | 1.54 | 0.62
| Analytical results for Co in OREAS 554 (Certified Value 0.761 wt.%) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 26 | 27 | 28 | 29 | 30 | wt.%
|  |  | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*AAS | AR*OES/MS | AR*OES | AR*MS | 3
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 3
| 1 | 1 | 0.7446 | 0.713 | 0.769 | 0.7451 | 0.786 | 0.75422 | > 0.5 | 0.826 | 0.6147 | 0.747 | 0.702 | 0.73 | 0.743 | 0.727 | 0.41093 | 0.82396 | 0.73434 | 0.76025 | > 0.5 | > 0.2 | 0.782 | 0.93521 | 0.78486 | 0.7592 | 0.831 | 1
| 1 | 2 | 0.7371 | 0.73201 | 0.774 | 0.7492 | 0.7756 | 0.76144 | > 0.5 | 0.832 | 0.6291 | 0.745 | 0.692 | 0.726 | 0.737 | 0.728 | 0.4 | 0.82578 | 0.71463 | 0.82289 | > 0.5 | > 0.2 | 0.828 | 0.93266 | 0.79535 | 0.7698 | 0.836 | 0
| 1 | 3 | 0.7591 | 0.71116 | 0.775 | 0.752 | 0.7904 | 0.76741 | > 0.5 | 0.802 | 0.6213 | 0.745 | 0.694 | 0.722 | 0.733 | 0.721 | 0.40642 | 0.83033 | 0.73778 | 0.74587 | > 0.5 | > 0.2 | 0.807 | 0.94387 | 0.77913 | 0.7457 | 0.833 | 16
| 1 | 4 | 0.7095 | 0.70038 | 0.761 | 0.7631 | 0.7885 | 0.74916 | > 0.5 | 0.818 | 0.6195 | 0.726 | 0.672 | 0.711 | 0.723 | 0.722 | 0.40669 | 0.83679 | 0.72628 | 0.76743 | > 0.5 | > 0.2 | 0.758 | 0.94152 | 0.79384 | 0.7629 | 0.869 | 0.76093
| 1 | 5 | 0.7346 | 0.69691 | 0.76 | 0.7564 | 0.7799 | 0.78094 | > 0.5 | 0.833 | 0.6123 | 0.707 | 0.666 | 0.725 | 0.704 | 0.735 | 0.41198 | 0.82356 | 0.74036 | 0.81267 | > 0.5 | > 0.2 | 0.804 | 0.93298 | 0.77683 | 0.7655 | 0.829 | 153
| 1 | 6 | 0.7417 | 0.71929 | 0.767 | 0.7629 | 0.7852 | 0.79265 | > 0.5 | 0.818 | 0.6281 | 0.729 | 0.672 | 0.719 | 0.754 | 0.728 | 0.40464 | 0.83174 | 0.73139 | 0.81191 | > 0.5 | > 0.2 | 0.769 | 0.78465 | 0.76596 | 0.7495 | 0.848
| Mean |  | 0.73777 | 0.71213 | 0.76767 | 0.75478 | 0.78427 | 0.76764 |  | 0.8215 | 0.62083 | 0.73317 | 0.683 | 0.72217 | 0.73233 | 0.72683 | 0.40678 | 0.82869 | 0.7308 | 0.78684 |  |  | 0.79133 | 0.91182 | 0.78266 | 0.75877 | 0.841
| Median |  | 0.7394 | 0.71208 | 0.768 | 0.7542 | 0.7856 | 0.76443 |  | 0.822 | 0.6204 | 0.737 | 0.682 | 0.7235 | 0.735 | 0.7275 | 0.40655 | 0.82806 | 0.73286 | 0.78967 |  |  | 0.793 | 0.9341 | 0.78199 | 0.76105 | 0.8345
| Std Dev. |  | 0.01629 | 0.01279 | 0.006314 | 0.0073532 | 0.0055424 | 0.016508 |  | 0.011554 | 0.0068348 | 0.015651 | 0.014792 | 0.0066156 | 0.017294 | 0.0050365 | 0.0043577 | 0.0051923 | 0.0093262 | 0.032734 |  |  | 0.026258 | 0.062466 | 0.011101 | 0.009391 | 0.015271
| Rel.Std.Dev. |  | 2.208% | 1.796% | 0.8225% | 0.9742% | 0.7067% | 2.15% |  | 1.406% | 1.101% | 2.135% | 2.166% | 0.9161% | 2.361% | 0.6929% | 1.071% | 0.6266% | 1.276% | 4.16% |  |  | 3.318% | 6.851% | 1.418% | 1.238% | 1.816%
| PDM3 |  | -3.044% | -6.414% | 0.8849% | -0.8082% | 3.066% | 0.8809% |  | 7.96% | -18.41% | -3.649% | -10.24% | -5.095% | -3.759% | -4.481% | -46.54% | 8.905% | -3.96% | 3.404% |  |  | 3.995% | 19.83% | 2.855% | -0.2847% | 10.52%
| Z-Score (Absolute) |  | 0.43 | 1 | 0.24 | 0.04 | 0.62 | 0.24 | NA | 1.45 | 3.06 | 0.53 | 1.66 | 0.78 | 0.55 | 0.67 | 7.87 | 1.62 | 0.58 | 0.68 | NA | NA | 0.78 | 3.48 | 0.58 | 0.04 | 1.89
| Analytical results for Cr in OREAS 554 (Certified Value 15.5 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 1
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g |  |  |  |  | 1
| 1 | 1 | 15 | 16 | 40 | 18 | 15 | 18.1 | 18 | 13.3 | 17 | 15 | 16 | 15 | 15 | 15 | 11 | 11.1697 | 12 | 16 | 21.9 | 17 | 14 |  |  |  |  | 1
| 1 | 2 | 15 | 17 | 20 | 17 | 20 | 22.9 | 17 | 13.7 | 16 | 14 | 17 | 15 | 16 | 15 | 9.8 | 13.7726 | 12 | 16 | 14.7 | 16 | 15 |  |  |  |  | 0
| 1 | 3 | 15 | 17 | 10 | 17 | 20 | 16.4 | 17 | 13.4 | 16 | 15 | 14 | 16 | 15 | 15 | 10.5 | 10.8526 | 21 | 16 | 15.3 | 16 | 15 |  |  |  |  | 16
| 1 | 4 | 14 | 16 | 10 | 17 | 15 | 17.1 | 17 | 13.3 | 16 | 15 | 17 | 14 | 14 | 20 | 14.9 | 10.9012 | 19 | 16 | 15.5 | 19 | 14 |  |  |  |  | 15.5232
| 1 | 5 | 14 | 16 | 10 | 17 | 20 | 15.7 | 18 | 13.7 | 16 | 17 | 16 | 15 | 14 | 14 | 10.5 | 11.4698 | 14 | 16 | 13.7 | 17 | 16 |  |  |  |  | 154
| 1 | 6 | 14 | 17 | 10 | 18 | 15 | 16.7 | 17 | 13.7 | 16 | 14 | 20 | 15 | 15 | 14 | 9.9 | 12.3597 | 26 | 15 | 14.5 | 15 | 16
| Mean |  | 14.5 | 16.5 | 16.6667 | 17.3333 | 17.5 | 17.8167 | 17.3333 | 13.5167 | 16.1667 | 15 | 16.6667 | 15 | 14.8333 | 15.5 | 11.1 | 11.7543 | 17.3333 | 15.8333 | 15.9333 | 16.6667 | 15
| Median |  | 14.5 | 16.5 | 10 | 17 | 17.5 | 16.9 | 17 | 13.55 | 16 | 15 | 16.5 | 15 | 15 | 15 | 10.5 | 11.3198 | 16.5 | 16 | 15 | 16.5 | 15
| Std Dev. |  | 0.54772 | 0.54772 | 12.1106 | 0.5164 | 2.7386 | 2.6141 | 0.5164 | 0.20412 | 0.40825 | 1.0954 | 1.9664 | 0.63246 | 0.75277 | 2.2583 | 1.9131 | 1.1317 | 5.6451 | 0.40825 | 2.9918 | 1.3663 | 0.89443
| Rel.Std.Dev. |  | 3.777% | 3.32% | 72.66% | 2.979% | 15.65% | 14.67% | 2.979% | 1.51% | 2.525% | 7.303% | 11.8% | 4.216% | 5.075% | 14.57% | 17.24% | 9.628% | 32.57% | 2.578% | 18.78% | 8.198% | 5.963%
| PDM3 |  | -6.591% | 6.293% | 7.366% | 11.66% | 12.73% | 14.77% | 11.66% | -12.93% | 4.145% | -3.37% | 7.366% | -3.37% | -4.444% | -0.1494% | -28.49% | -24.28% | 11.66% | 1.998% | 2.642% | 7.366% | -3.37%
| Z-Score (Absolute) |  | 0.96 | 0.43 | NA | 1.01 | 1.13 | 1.35 | 1.01 | 1.65 | 0.2 | 0.62 | 0.55 | 0.62 | 0.73 | 0.27 | 3.34 | 2.88 | 1.01 | 0.03 | 0.03 | 0.55 | 0.62
Results from laboratory 4 were removed due to their 10 ppm reading resolution.
| Analytical results for Cs in OREAS 554 (Certified Value 0.57 ppm) | [26 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [7 empty] | 2
| 1 | 1 | 0.71 | 0.6 | 0.69 | 0.5 | 0.65 | 0.66 | 0.46 | 0.51 | 0.49 | 0.56 | 0.54 | 0.4 | 0.50638 | 0.49 | 0.86 | 0.72 | 0.53 | 0.42 | [7 empty] | 1
| 1 | 2 | 0.72 | 0.6 | 0.72 | 0.5 | 0.66 | 0.63 | 0.47 | 0.48 | 0.5 | 0.6 | 0.51 | 0.5 | 0.63627 | 0.48 | 0.87 | 0.74 | 0.5 | 0.43 | [7 empty] | 0
| 1 | 3 | 0.68 | 0.6 | 0.73 | 0.5 | 0.64 | 0.7 | 0.48 | 0.5 | 0.49 | 0.6 | 0.51 | 0.5 | 0.49753 | 0.49 | 0.88 | 0.69 | 0.52 | 0.41 | [7 empty] | 16
| 1 | 4 | 0.65 | 0.6 | 0.72 | 0.5 | 0.66 | 0.64 | 0.46 | 0.51 | 0.47 | 0.6 | 0.45 | 0.5 | 0.52941 | 0.5 | 0.87 | 0.73 | 0.52 | 0.42 | [7 empty] | 0.56649
| 1 | 5 | 0.67 | 0.6 | 0.71 | 0.5 | 0.68 | 0.62 | 0.47 | 0.5 | 0.47 | 0.58 | 0.46 | 0.5 | 0.62146 | 0.51 | 0.86 | 0.73 | 0.52 | 0.43 | [7 empty] | 155
| 1 | 6 | 0.69 | 0.6 | 0.7 | 0.5 | 0.62 | 0.67 | 0.47 | 0.49 | 0.49 | 0.59 | 0.44 | 0.4 | 0.54429 | 0.49 | 0.87 | 0.68 | 0.52 | 0.41
| Mean |  | 0.68667 | 0.6 | 0.71167 | 0.5 | 0.65167 | 0.65333 | 0.46833 | 0.49833 | 0.485 | 0.58833 | 0.485 | 0.46667 | 0.55589 | 0.49333 | 0.86833 | 0.715 | 0.51833 | 0.42
| Median |  | 0.685 | 0.6 | 0.715 | 0.5 | 0.655 | 0.65 | 0.47 | 0.5 | 0.49 | 0.595 | 0.485 | 0.5 | 0.53685 | 0.49 | 0.87 | 0.725 | 0.52 | 0.42
| Std Dev. |  | 0.02582 | 0 | 0.01472 | 0 | 0.020412 | 0.029439 | 0.0075277 | 0.01169 | 0.012247 | 0.016021 | 0.040373 | 0.05164 | 0.059082 | 0.010328 | 0.0075277 | 0.02429 | 0.0098319 | 0.0089443
| Rel.Std.Dev. |  | 3.76% | 0% | 2.068% | 0% | 3.132% | 4.506% | 1.607% | 2.346% | 2.525% | 2.723% | 8.324% | 11.07% | 10.63% | 2.094% | 0.8669% | 3.397% | 1.897% | 2.13%
| PDM3 |  | 21.21% | 5.915% | 25.63% | -11.74% | 15.04% | 15.33% | -17.33% | -12.03% | -14.39% | 3.856% | -14.39% | -17.62% | -1.872% | -12.91% | 53.28% | 26.22% | -8.501% | -25.86%
| Z-Score (Absolute) |  | 1.01 | NA | 1.2 | NA | 0.74 | 0.75 | 0.67 | 0.44 | 0.55 | 0.25 | 0.55 | NA | 0 | 0.48 | 2.41 | 1.23 | 0.29 | 1.05
Results from laboratories 4, 6 & 18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 554 (Certified Value 1.58 wt.%) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 26 | 27 | 28 | 29 | 30 | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*AAS | AR*OES/MS | AR*AAS | AR*OES | 2
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 3
| 1 | 1 | 1.48 | 1.5202 | 1.55 | 1.6069 | 1.5327 | 1.523 | > 1 | 1.67 | 1.585 | 1.565 | 1.605 | 1.56 | 1.555 | 1.3153 | 1.621 | 1.61 | 1.5611 | 1.5231 | 1.49 | > 1 | 1.619 | 1.6816 | > 1 | 1.59 | 1.63 | 1
| 1 | 2 | 1.5 | 1.5518 | 1.54 | 1.6095 | 1.5587 | 1.5005 | > 1 | 1.67 | 1.55 | 1.54 | 1.63 | 1.57 | 1.595 | 1.3368 | 1.616 | 1.67 | 1.5248 | 1.5323 | 1.56 | > 1 | 1.671 | 1.7285 | > 1 | 1.64 | 1.63 | 0
| 1 | 3 | 1.5 | 1.5262 | 1.53 | 1.612 | 1.5624 | 1.5358 | > 1 | 1.67 | 1.605 | 1.545 | 1.58 | 1.555 | 1.595 | 1.3718 | 1.632 | 1.64 | 1.7045 | 1.5466 | 1.56 | > 1 | 1.612 | 1.7316 | > 1 | 1.61 | 1.62 | 16
| 1 | 4 | 1.49 | 1.503 | 1.5 | 1.6111 | 1.5396 | 1.5384 | > 1 | 1.65 | 1.595 | 1.545 | 1.625 | 1.535 | 1.515 | 1.3395 | 1.627 | 1.65 | 1.5357 | 1.5293 | 1.53 | > 1 | 1.575 | 1.5786 | > 1 | 1.58 | 1.62 | 1.5785
| 1 | 5 | 1.52 | 1.4834 | 1.55 | 1.6451 | 1.5793 | 1.552 | > 1 | 1.66 | 1.585 | 1.55 | 1.605 | 1.57 | 1.51 | 1.3505 | 1.612 | 1.63 | 1.5835 | 1.5683 | 1.53 | > 1 | 1.602 | 1.7252 | > 1 | 1.59 | 1.61 | 156
| 1 | 6 | 1.51 | 1.5252 | 1.57 | 1.6497 | 1.5585 | 1.5503 | > 1 | 1.68 | 1.585 | 1.535 | 1.595 | 1.57 | 1.515 | 1.3258 | 1.629 | 1.61 | 1.5793 | 1.5361 | 1.54 | > 1 | 1.585 | 1.5181 | > 1 | 1.58 | 1.6
| Mean |  | 1.5 | 1.5183 | 1.54 | 1.6224 | 1.5552 | 1.5333 |  | 1.6667 | 1.5842 | 1.5467 | 1.6067 | 1.56 | 1.5475 | 1.3399 | 1.6228 | 1.635 | 1.5815 | 1.5393 | 1.535 |  | 1.6107 | 1.6606 |  | 1.5983 | 1.6183
| Median |  | 1.5 | 1.5227 | 1.545 | 1.6116 | 1.5586 | 1.5371 |  | 1.67 | 1.585 | 1.545 | 1.605 | 1.565 | 1.535 | 1.3381 | 1.624 | 1.635 | 1.5702 | 1.5342 | 1.535 |  | 1.607 | 1.7034 |  | 1.59 | 1.62
| Std Dev. |  | 0.014142 | 0.023192 | 0.023664 | 0.019509 | 0.016765 | 0.019265 |  | 0.010328 | 0.018552 | 0.010328 | 0.018619 | 0.013784 | 0.040218 | 0.019707 | 0.0078337 | 0.023452 | 0.064577 | 0.016229 | 0.025884 |  | 0.033815 | 0.090862 |  | 0.023166 | 0.01169
| Rel.Std.Dev. |  | 0.9428% | 1.527% | 1.537% | 1.203% | 1.078% | 1.256% |  | 0.6197% | 1.171% | 0.6678% | 1.159% | 0.8836% | 2.599% | 1.471% | 0.4827% | 1.434% | 4.083% | 1.054% | 1.686% |  | 2.099% | 5.472% |  | 1.449% | 0.7224%
| PDM3 |  | -4.976% | -3.816% | -2.442% | 2.777% | -1.479% | -2.865% |  | 5.582% | 0.356% | -2.02% | 1.781% | -1.175% | -1.967% | -15.12% | 2.806% | 3.576% | 0.1862% | -2.488% | -2.759% |  | 2.035% | 5.198% |  | 1.253% | 2.52%
| Z-Score (Absolute) |  | 1.3 | 0.96 | 0.57 | 0.95 | 0.29 | 0.69 | NA | 1.76 | 0.25 | 0.44 | 0.66 | 0.2 | 0.43 | 4.24 | 0.96 | 1.18 | 0.2 | 0.58 | 0.66 | NA | 0.73 | 1.65 | NA | 0.51 | 0.88
| Analytical results for Dy in OREAS 554 (Certified Value 1.34 ppm) | [26 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 30 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.5g | 2
| 1 | 1 | 01: 1.472; 02: 1.29; 04: 1.3; 05: 1.6; 06: 1.18 | 1
| 1 | 2 | 01: 1.471; 02: 1.32; 04: 1.2; 05: 1.5; 06: 1.17 | 0
| 1 | 3 | 01: 1.455; 02: 1.26; 04: 1.3; 05: 1.6; 06: 1.19 | 16
| 1 | 4 | 01: 1.398; 02: 1.25; 04: 1.3; 05: 1.5; 06: 1.2 | 1.343
| 1 | 5 | 01: 1.405; 02: 1.23; 04: 1.2; 05: 1.6; 06: 1.17 | 157
| 1 | 6 | 01: 1.488; 02: 1.29; 04: 1.3; 05: 1.5; 06: 1.15
| Mean |  | 01: 1.4482; 02: 1.2733; 04: 1.2667; 05: 1.55; 06: 1.1767
| Median |  | 01: 1.463; 02: 1.275; 04: 1.3; 05: 1.55; 06: 1.175
| Std Dev. |  | 01: 0.03769; 02: 0.03266; 04: 0.05164; 05: 0.054772; 06: 0.017512
| Rel.Std.Dev. |  | 01: 2.603%; 02: 2.565%; 04: 4.077%; 05: 3.534%; 06: 1.488%
| PDM3 |  | 01: 7.833%; 02: -5.185%; 04: -5.681%; 05: 15.42%; 06: -12.38%
| Z-Score (Absolute) |  | 01: 1.22; 02: 0; 04: 0.05; 05: 1.93; 06: 0.67
| Analytical results for Er in OREAS 554 (Certified Value 0.57 ppm) | [26 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 30 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.5g | 3
| 1 | 1 | 01: 0.631; 02: 0.57; 04: 0.6; 05: 0.6; 06: 0.53 | 1
| 1 | 2 | 01: 0.64; 02: 0.57; 04: 0.55; 05: 0.6; 06: 0.51 | 0
| 1 | 3 | 01: 0.616; 02: 0.54; 04: 0.55; 05: 0.6; 06: 0.55 | 16
| 1 | 4 | 01: 0.595; 02: 0.54; 04: 0.6; 05: 0.5; 06: 0.63 | 0.57173
| 1 | 5 | 01: 0.613; 02: 0.54; 04: 0.55; 05: 0.6; 06: 0.51 | 158
| 1 | 6 | 01: 0.619; 02: 0.56; 04: 0.6; 05: 0.6; 06: 0.54
| Mean |  | 01: 0.619; 02: 0.55333; 04: 0.575; 05: 0.58333; 06: 0.545
| Median |  | 01: 0.6175; 02: 0.55; 04: 0.575; 05: 0.6; 06: 0.535
| Std Dev. |  | 01: 0.015531; 02: 0.015055; 04: 0.027386; 05: 0.040825; 06: 0.044609
| Rel.Std.Dev. |  | 01: 2.509%; 02: 2.721%; 04: 4.763%; 05: 6.999%; 06: 8.185%
| PDM3 |  | 01: 8.267%; 02: -3.218%; 04: 0.5714%; 05: 2.029%; 06: -4.676%
| Z-Score (Absolute) |  | 01: 1.37; 02: 0.67; 04: 0; 05: 0.26; 06: 0.93
| Analytical results for Eu in OREAS 554 (Certified Value 0.45 ppm) | [26 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 30 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.5g | 2
| 1 | 1 | 01: 0.519; 02: 0.404; 04: 0.5; 05: 0.5; 06: 0.33 | 1
| 1 | 2 | 01: 0.549; 02: 0.386; 04: 0.5; 05: 0.5; 06: 0.33 | 0
| 1 | 3 | 01: 0.536; 02: 0.4; 04: 0.5; 05: 0.5; 06: 0.32 | 16
| 1 | 4 | 01: 0.508; 02: 0.396; 04: 0.5; 05: 0.5; 06: 0.34 | 0.4545
| 1 | 5 | 01: 0.519; 02: 0.388; 04: 0.5; 05: 0.6; 06: 0.31 | 159
| 1 | 6 | 01: 0.538; 02: 0.412; 04: 0.5; 05: 0.5; 06: 0.35
| Mean |  | 01: 0.52817; 02: 0.39767; 04: 0.5; 05: 0.51667; 06: 0.33
| Median |  | 01: 0.5275; 02: 0.398; 04: 0.5; 05: 0.5; 06: 0.33
| Std Dev. |  | 01: 0.015276; 02: 0.0098319; 04: 0; 05: 0.040825; 06: 0.014142
| Rel.Std.Dev. |  | 01: 2.892%; 02: 2.472%; 04: 0%; 05: 7.902%; 06: 4.285%
| PDM3 |  | 01: 16.21%; 02: -12.5%; 04: 10.01%; 05: 13.68%; 06: -27.39%
| Z-Score (Absolute) |  | 01: 0.67; 02: 2.45; 04: 0; 05: 0.4; 06: 4.07
| Analytical results for Fe in OREAS 554 (Certified Value 1.64 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 27 | 28 | 29 |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*AAS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  | 3
| 1 | 1 | 1.59 | 1.65 | 1.68 | 1.63 | 1.711 | 1.62 | 1.65 | 1.67 | 1.7 | 1.63 | 1.63 | 1.67 | 1.67 | 1.63 | 1.5202 | 1.6956 | 1.63 | 1.7 | 1.56 | 1.59 | 1.7176 | 1.67 | 1.68 |  |  | 1
| 1 | 2 | 1.63 | 1.69 | 1.67 | 1.63 | 1.653 | 1.6 | 1.62 | 1.67 | 1.73 | 1.63 | 1.61 | 1.67 | 1.67 | 1.62 | 1.5359 | 1.6201 | 1.66 | 1.67 | 1.57 | 1.6 | 1.7118 | 1.71 | 1.66 |  |  | 0
| 1 | 3 | 1.6 | 1.67 | 1.61 | 1.68 | 1.673 | 1.63 | 1.55 | 1.72 | 1.73 | 1.65 | 1.62 | 1.67 | 1.66 | 1.61 | 1.5788 | 1.6591 | 1.58 | 1.66 | 1.57 | 1.6 | 1.7228 | 1.65 | 1.63 |  |  | 16
| 1 | 4 | 1.53 | 1.62 | 1.69 | 1.65 | 1.703 | 1.62 | 1.61 | 1.64 | 1.72 | 1.62 | 1.59 | 1.63 | 1.63 | 1.57 | 1.5347 | 1.698 | 1.63 | 1.64 | 1.56 | 1.6 | 1.7192 | 1.69 | 1.64 |  |  | 1.6391
| 1 | 5 | 1.57 | 1.6 | 1.67 | 1.67 | 1.6635 | 1.63 | 1.65 | 1.66 | 1.75 | 1.6 | 1.57 | 1.63 | 1.61 | 1.58 | 1.5533 | 1.6309 | 1.53 | 1.7 | 1.56 | 1.52 | 1.7326 | 1.67 | 1.65 |  |  | 160
| 1 | 6 | 1.58 | 1.65 | 1.69 | 1.62 | 1.7115 | 1.64 | 1.58 | 1.7 | 1.74 | 1.63 | 1.6 | 1.66 | 1.67 | 1.58 | 1.55 | 1.6146 | 1.64 | 1.7 | 1.56 | 1.53 | 1.625 | 1.63 | 1.65
| Mean |  | 1.5833 | 1.6467 | 1.6683 | 1.6467 | 1.6858 | 1.6233 | 1.61 | 1.6767 | 1.7283 | 1.6267 | 1.6033 | 1.655 | 1.6517 | 1.5983 | 1.5455 | 1.6531 | 1.6117 | 1.6783 | 1.5633 | 1.5733 | 1.7048 | 1.67 | 1.6517
| Median |  | 1.585 | 1.65 | 1.675 | 1.64 | 1.688 | 1.625 | 1.615 | 1.67 | 1.73 | 1.63 | 1.605 | 1.665 | 1.665 | 1.595 | 1.543 | 1.645 | 1.63 | 1.685 | 1.56 | 1.595 | 1.7184 | 1.67 | 1.65
| Std Dev. |  | 0.033267 | 0.03266 | 0.029944 | 0.024221 | 0.025801 | 0.013663 | 0.039497 | 0.028752 | 0.017224 | 0.01633 | 0.021602 | 0.019748 | 0.025626 | 0.024833 | 0.020193 | 0.037208 | 0.047924 | 0.025626 | 0.005164 | 0.037771 | 0.039709 | 0.028284 | 0.017224
| Rel.Std.Dev. |  | 2.101% | 1.983% | 1.795% | 1.471% | 1.53% | 0.8416% | 2.453% | 1.715% | 0.9966% | 1.004% | 1.347% | 1.193% | 1.551% | 1.554% | 1.307% | 2.251% | 2.974% | 1.527% | 0.3303% | 2.401% | 2.329% | 1.694% | 1.043%
| PDM3 |  | -3.404% | 0.4602% | 1.782% | 0.4602% | 2.85% | -0.9633% | -1.777% | 2.29% | 5.443% | -0.7599% | -2.183% | 0.9686% | 0.7653% | -2.489% | -5.713% | 0.8499% | -1.675% | 2.392% | -4.624% | -4.014% | 4.009% | 1.884% | 0.7653%
| Z-Score (Absolute) |  | 1.35 | 0 | 0.46 | 0 | 0.83 | 0.5 | 0.78 | 0.64 | 1.74 | 0.43 | 0.92 | 0.18 | 0.11 | 1.03 | 2.15 | 0.14 | 0.75 | 0.67 | 1.77 | 1.56 | 1.24 | 0.5 | 0.11
| Analytical results for Ga in OREAS 554 (Certified Value 1.76 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  | 2
| 1 | 1 | 1.8 | 1.45 | 1.8 | 1.7 | 2 | < 5 | 1.78 | 2.2 | 2 | 1.82 | 1.92 | 1.89 | 1.92 | 2.06 | 1.28 | 1.5564 | 1.4 | 3.4 | 1.7 | 2.25 | 1.8 | 1.08 |  |  |  | 1
| 1 | 2 | 1.8 | 1.62 | 1.8 | 1.7 | 2 | < 5 | 1.52 | 2.3 | 3 | 1.81 | 1.88 | 1.87 | 1.95 | 2.12 | 1.09 | 1.6114 | 1.3 | 3.4 | 1.9 | 2.32 | 1.9 | 1.04 |  |  |  | 0
| 1 | 3 | 1.6 | 1.64 | 1.6 | 1.8 | 2 | < 5 | 1.64 | 2.3 | 2 | 1.76 | 1.92 | 1.87 | 1.93 | 2.03 | 1.39 | 1.4459 | 1.4 | 3.3 | 1.7 | 2.2 | 1.8 | 1.06 |  |  |  | 16
| 1 | 4 | 1.6 | 1.37 | 1.6 | 1.7 | 2 | < 5 | 1.71 | 2.3 | 2 | 1.72 | 2.03 | 1.79 | 1.95 | 1.79 | 1.33 | 1.5405 | 1.4 | 3.3 | 1.9 | 2.27 | 1.9 | 1.01 |  |  |  | 1.7642
| 1 | 5 | 1.6 | 1.53 | 1.6 | 1.8 | 2 | < 5 | 1.83 | 2.2 | 2 | 1.71 | 2 | 1.82 | 1.91 | 1.75 | 1.25 | 1.433 | 1.4 | 3.2 | 1.7 | 2.21 | 1.9 | 1.04 |  |  |  | 161
| 1 | 6 | 1.5 | 1.54 | 1.7 | 1.7 | 2 | < 5 | 1.76 | 2.3 | 3 | 1.78 | 1.9 | 1.83 | 1.95 | 1.82 | 1.3 | 1.6099 | 1.4 | 3.3 | 1.6 | 2.07 | 1.8 | 1.09
| Mean |  | 1.65 | 1.525 | 1.6833 | 1.7333 | 2 |  | 1.7067 | 2.2667 | 2.3333 | 1.7667 | 1.9417 | 1.845 | 1.935 | 1.9283 | 1.2733 | 1.5328 | 1.3833 | 3.3167 | 1.75 | 2.22 | 1.85 | 1.0533
| Median |  | 1.6 | 1.535 | 1.65 | 1.7 | 2 |  | 1.735 | 2.3 | 2 | 1.77 | 1.92 | 1.85 | 1.94 | 1.925 | 1.29 | 1.5484 | 1.4 | 3.3 | 1.7 | 2.23 | 1.85 | 1.05
| Std Dev. |  | 0.12247 | 0.10213 | 0.098319 | 0.05164 | 0 |  | 0.11201 | 0.05164 | 0.5164 | 0.045461 | 0.05947 | 0.037815 | 0.017607 | 0.15943 | 0.10172 | 0.077779 | 0.040825 | 0.075277 | 0.12247 | 0.085323 | 0.054772 | 0.029439
| Rel.Std.Dev. |  | 7.423% | 6.697% | 5.841% | 2.979% | 0% |  | 6.563% | 2.278% | 22.13% | 2.573% | 3.063% | 2.05% | 0.9099% | 8.268% | 7.988% | 5.074% | 2.951% | 2.27% | 6.999% | 3.843% | 2.961% | 2.795%
| PDM3 |  | -6.472% | -13.56% | -4.583% | -1.749% | 13.37% |  | -3.26% | 28.48% | 32.26% | 0.1406% | 10.06% | 4.581% | 9.682% | 9.304% | -27.82% | -13.11% | -21.59% | 88% | -0.8041% | 25.84% | 4.864% | -40.29%
| Z-Score (Absolute) |  | 0.41 | 0.87 | 0.28 | 0.09 | NA | 2.78 | 0.19 | 1.9 | NA | 0.03 | 0.69 | 0.32 | 0.66 | 0.64 | 1.82 | 0.84 | 1.4 | 5.84 | 0.03 | 1.73 | 0.34 | 2.64
Results from laboratories 6 & 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 554 (Certified Value 2.18 ppm) | [26 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 30 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS |  |  |  | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.5g |  |  |  | 2
| 1 | 1 | 01: 2.429; 02: 1.93; 04: 2.5; 05: 2.7; 06: 1.51 |  |  |  | 1
| 1 | 2 | 01: 2.5; 02: 1.97; 04: 2.4; 05: 2.7; 06: 1.41 |  |  |  | 0
| 1 | 3 | 01: 2.438; 02: 1.85; 04: 2.4; 05: 2.6; 06: 1.5 |  |  |  | 16
| 1 | 4 | 01: 2.308; 02: 1.88; 04: 2.5; 05: 2.6; 06: 1.53 |  |  |  | 2.1755
| 1 | 5 | 01: 2.395; 02: 1.82; 04: 2.3; 05: 2.7; 06: 1.39 |  |  |  | 162
| 1 | 6 | 01: 2.446; 02: 1.84; 04: 2.5; 05: 2.7; 06: 1.52
| Mean |  | 01: 2.4193; 02: 1.8817; 04: 2.4333; 05: 2.6667; 06: 1.4767
| Median |  | 01: 2.4335; 02: 1.865; 04: 2.45; 05: 2.7; 06: 1.505
| Std Dev. |  | 01: 0.064261; 02: 0.057764; 04: 0.08165; 05: 0.05164; 06: 0.060553
| Rel.Std.Dev. |  | 01: 2.656%; 02: 3.07%; 04: 3.355%; 05: 1.936%; 06: 4.101%
| PDM3 |  | 01: 11.21%; 02: -13.51%; 04: 11.85%; 05: 22.58%; 06: -32.12%
| Z-Score (Absolute) |  | 01: 0; 02: 1.47; 04: 0.04; 05: 0.67; 06: 2.57
| Analytical results for Ge in OREAS 554 (Indicative Value 0.045 ppm) | [26 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 06 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 28 | 29 | 30 | [11 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 0.25g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | 0.05 | 0.06 | < 0.05 | 0.02 | 0.052755 | < 0.1 | 0.07 | < 0.1 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 0.1 | 0.05 | 0.05 | 0.05 | 0.06 | < 0.05 | 0.01 | 0.061243 | < 0.1 | 0.07 | < 0.1 | < 0.01 | [11 empty] | 0
| 1 | 3 | < 0.1 | < 0.1 | < 0.1 | 0.05 | < 0.05 | 0.06 | 0.07 | < 0.05 | 0.02 | 0.043043 | < 0.1 | 0.07 | < 0.1 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | 0.05 | 0.05 | 0.06 | 0.05 | 0.02 | 0.0505 | < 0.1 | 0.06 | < 0.1 | < 0.01 | [11 empty] | 0.045334
| 1 | 5 | < 0.1 | < 0.1 | < 0.1 | 0.05 | 0.05 | 0.06 | 0.07 | < 0.05 | 0.01 | 0.069384 | < 0.1 | 0.07 | < 0.1 | < 0.01 | [11 empty] | 28
| 1 | 6 | < 0.1 | < 0.1 | < 0.1 | < 0.05 | 0.06 | 0.06 | 0.06 | < 0.05 | 0.01 | 0.054103 | < 0.1 | 0.07 | < 0.1 | < 0.01
| Mean |  | 05: 0.05; 06: 0.0525; 09: 0.055; 10: 0.063333; 11: 0.05; 12: 0.015; 13: 0.055171; 15: 0.068333
| Median |  | 05: 0.05; 06: 0.05; 09: 0.055; 10: 0.06; 11: 0.05; 12: 0.015; 13: 0.053429; 15: 0.07
| Std Dev. |  | 05: 8.4984e-18; 06: 0.005; 09: 0.0054772; 10: 0.005164; 12: 0.0054772; 13: 0.0091077; 15: 0.0040825
| Rel.Std.Dev. |  | 05: 1.7e-14%; 06: 9.524%; 09: 9.959%; 10: 8.154%; 12: 36.51%; 13: 16.51%; 15: 5.974%
| PDM3 |  | 05: 10.29%; 06: 15.81%; 09: 21.32%; 10: 39.7%; 11: 10.29%; 12: -66.91%; 13: 21.7%; 15: 50.73%
| Z-Score (Absolute) |  | 0 | 0 | 0 | 1.42 | 0.76 | 0.57 | 1.52 | 2.37 | 3.99 | 0.59 | 0 | 2.09 | 0 | 5.13
| Analytical results for Hf in OREAS 554 (Certified Value 0.27 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 05 | 06 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | 30 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 2
|  |  | 25g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [8 empty] | 2
| 1 | 1 | 0.24 | 0.25 | < 0.5 | 0.2 | 0.35 | 0.26 | 0.25 | 0.24 | 0.25 | 0.27 | 0.46 | 0.42754 | 0.27 | 0.26 | 0.29 | 0.28 | 0.23 | [8 empty] | 1
| 1 | 2 | 0.25 | 0.24 | < 0.5 | 0.2 | 0.31 | 0.24 | 0.26 | 0.24 | 0.26 | 0.27 | 0.45 | 0.48753 | 0.28 | 0.28 | 0.3 | 0.3 | 0.22 | [8 empty] | 0
| 1 | 3 | 0.23 | 0.26 | < 0.5 | 0.2 | 0.35 | 0.26 | 0.25 | 0.24 | 0.28 | 0.28 | 0.52 | 0.45557 | 0.28 | 0.27 | 0.28 | 0.31 | 0.23 | [8 empty] | 16
| 1 | 4 | 0.23 | 0.24 | < 0.5 | 0.2 | 0.34 | 0.26 | 0.25 | 0.24 | 0.27 | 0.28 | 0.47 | 0.46903 | 0.3 | 0.25 | 0.29 | 0.31 | 0.23 | [8 empty] | 0.26654
| 1 | 5 | 0.24 | 0.25 | < 0.5 | 0.2 | 0.33 | 0.24 | 0.25 | 0.24 | 0.29 | 0.28 | 0.45 | 0.43715 | 0.26 | 0.25 | 0.28 | 0.3 | 0.22 | [8 empty] | 163
| 1 | 6 | 0.24 | 0.26 | < 0.5 | 0.2 | 0.35 | 0.25 | 0.26 | 0.23 | 0.27 | 0.27 | 0.45 | 0.45468 | 0.28 | 0.27 | 0.27 | 0.29 | 0.22
| Mean |  | 0.23833 | 0.25 |  | 0.2 | 0.33833 | 0.25167 | 0.25333 | 0.23833 | 0.27 | 0.275 | 0.46667 | 0.45525 | 0.27833 | 0.26333 | 0.285 | 0.29833 | 0.225
| Median |  | 0.24 | 0.25 |  | 0.2 | 0.345 | 0.255 | 0.25 | 0.24 | 0.27 | 0.275 | 0.455 | 0.45512 | 0.28 | 0.265 | 0.285 | 0.3 | 0.225
| Std Dev. |  | 0.0075277 | 0.0089443 |  | 3.0405e-17 | 0.016021 | 0.0098319 | 0.005164 | 0.0040825 | 0.014142 | 0.0054772 | 0.027325 | 0.021579 | 0.013292 | 0.012111 | 0.010488 | 0.01169 | 0.0054772
| Rel.Std.Dev. |  | 3.158% | 3.578% |  | 1.52e-14% | 4.735% | 3.907% | 2.038% | 1.713% | 5.238% | 1.992% | 5.855% | 4.74% | 4.775% | 4.599% | 3.68% | 3.919% | 2.434%
| PDM3 |  | -10.58% | -6.205% |  | -24.96% | 26.94% | -5.58% | -4.954% | -10.58% | 1.299% | 3.175% | 75.08% | 70.8% | 4.425% | -1.203% | 6.926% | 11.93% | -15.58%
| Z-Score (Absolute) |  | 1.09 | 0.64 | 0.64 | NA | 2.76 | 0.58 | 0.51 | 1.09 | 0.13 | 0.32 | 7.71 | 7.27 | 0.45 | 0.13 | 0.71 | 1.22 | 1.61
Results from laboratory 10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 554 (Certified Value 0.019 ppm) | [26 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 24 | 28 | 29 | 30 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.01 | < 0.1 | 0.02 | < 0.05 | 0.01 | < 0.05 | 0.03 | 0.56 | 0.02 | 0.02 | 0.03 | 0.02 | 0.03 | 0.016 | < 0.1 | < 0.01 | < 0.005 | 0.021 | < 0.01 | 0.05 | [5 empty] | 1
| 1 | 2 | 0.02 | < 0.1 | 0.025 | < 0.05 | < 0.01 | < 0.05 | 0.07 | 0.53 | 0.01 | 0.02 | 0.03 | 0.02 | 0.02 | 0.014 | < 0.1 | 0.02 | < 0.005 | 0.024 | < 0.01 | 0.04 | [5 empty] | 0
| 1 | 3 | 0.01 | < 0.1 | 0.015 | < 0.05 | 0.01 | < 0.05 | 0.05 | 0.54 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.018 | < 0.1 | < 0.01 | < 0.005 | 0.023 | < 0.01 | 0.04 | [5 empty] | 16
| 1 | 4 | 0.01 | < 0.1 | 0.02 | < 0.05 | 0.01 | < 0.05 | 0.05 | 0.49 | 0.02 | 0.03 | 0.02 | 0.02 | 0.03 | 0.016 | < 0.1 | < 0.01 | < 0.005 | 0.021 | < 0.01 | 0.04 | [5 empty] | 0.018636
| 1 | 5 | 0.01 | < 0.1 | 0.02 | < 0.05 | 0.01 | < 0.05 | 0.03 | 0.47 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.015 | < 0.1 | 0.01 | < 0.005 | 0.022 | < 0.01 | 0.04 | [5 empty] | 164
| 1 | 6 | < 0.01 | < 0.1 | 0.02 | < 0.05 | 0.01 | < 0.05 | 0.04 | 0.59 | 0.02 | 0.01 | 0.03 | 0.02 | 0.02 | 0.015 | < 0.1 | 0.02 | < 0.005 | 0.023 | < 0.01 | 0.04
| Mean |  | 0.012 |  | 0.02 |  | 0.01 |  | 0.045 | 0.53 | 0.018333 | 0.02 | 0.026667 | 0.02 | 0.023333 | 0.015667 |  | 0.016667 |  | 0.022333 |  | 0.041667
| Median |  | 0.01 |  | 0.02 |  | 0.01 |  | 0.045 | 0.535 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | 0.0155 |  | 0.02 |  | 0.0225 |  | 0.04
| Std Dev. |  | 0.0044721 |  | 0.0031623 |  | 0 |  | 0.015166 | 0.044272 | 0.0040825 | 0.0063246 | 0.005164 | 0 | 0.005164 | 0.0013663 |  | 0.0057735 |  | 0.0012111 |  | 0.0040825
| Rel.Std.Dev. |  | 37.27% |  | 15.81% |  | 0% |  | 33.7% | 8.353% | 22.27% | 31.62% | 19.36% | 0% | 22.13% | 8.721% |  | 34.64% |  | 5.423% |  | 9.798%
| PDM3 |  | -35.61% |  | 7.317% |  | -46.34% |  | 141.5% | 2744% | -1.626% | 7.317% | 43.09% | 7.317% | 25.2% | -15.93% |  | -10.57% |  | 19.84% |  | 123.6%
| Z-Score (Absolute) |  | 0.88 | 2.46 | 0.1 | 0.33 | 1.02 | 0.33 | 2.03 | 43.34 | 0.24 | 0.1 | 0.47 | 0.1 | 0.18 | 0.47 | 2.46 | 0.88 | 1.59 | 0.1 | 1.38 | 1.75
| Analytical results for Ho in OREAS 554 (Certified Value 0.22 ppm) | [26 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab |  |  |  | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 19; 09: 30 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*OES/MS; 09: AR*MS |  |  |  | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 01g; 09: 0.5g |  |  |  | 2
| 1 | 1 | 01: 0.2; 02: 0.21; 04: 0.25; 05: 0.2; 06: 0.24403; 09: 0.22 |  |  |  | 1
| 1 | 2 | 01: 0.2; 02: 0.215; 04: 0.25; 05: 0.3; 06: 0.24488; 09: 0.2 |  |  |  | 0
| 1 | 3 | 01: 0.2; 02: 0.215; 04: 0.25; 05: 0.2; 06: 0.22103; 09: 0.19 |  |  |  | 16
| 1 | 4 | 01: 0.2; 02: 0.21; 04: 0.25; 05: 0.2; 06: 0.23638; 09: 0.2 |  |  |  | 0.22235
| 1 | 5 | 01: 0.2; 02: 0.21; 04: 0.2; 05: 0.3; 06: 0.22317; 09: 0.2 |  |  |  | 165
| 1 | 6 | 01: 0.2; 02: 0.215; 04: 0.25; 05: 0.2; 06: 0.25007; 09: 0.2
| Mean |  | 01: 0.2; 02: 0.2125; 04: 0.24167; 05: 0.23333; 06: 0.23659; 09: 0.20167
| Median |  | 01: 0.2; 02: 0.2125; 04: 0.25; 05: 0.2; 06: 0.24021; 09: 0.2
| Std Dev. |  | 01: 3.0405e-17; 02: 0.0027386; 04: 0.020412; 05: 0.05164; 06: 0.012067; 09: 0.0098319
| Rel.Std.Dev. |  | 01: 1.52e-14%; 02: 1.289%; 04: 8.447%; 05: 22.13%; 06: 5.1%; 09: 4.875%
| PDM3 |  | 01: -10.05%; 02: -4.429%; 04: 8.688%; 05: 4.94%; 06: 6.406%; 09: -9.302%
| Z-Score (Absolute) |  | 01: 0.95; 02: 0.43; 04: 0.78; 05: 0.43; 06: 0.57; 09: 0.88
| Analytical results for In in OREAS 554 (Certified Value 0.063 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 22 | 23 | 28 | 29 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 25g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.25g | [7 empty] | 3
| 1 | 1 | 0.06 | 0.06 | < 0.1 | 0.07 | 0.06 | 15 | 0.063 | 0.068 | 0.063 | 0.062 | 0.065 | 0.04 | < 0.5 | 0.07 | 0.06127 | 0.07 | 0.066 | 0.06 | [7 empty] | 1
| 1 | 2 | 0.07 | 0.06 | < 0.1 | 0.07 | 0.06 | 15 | 0.063 | 0.057 | 0.063 | 0.063 | 0.067 | 0.04 | < 0.5 | 0.06 | 0.06322 | 0.07 | 0.067 | 0.06 | [7 empty] | 0
| 1 | 3 | 0.06 | 0.06 | < 0.1 | 0.05 | 0.1 | 15 | 0.066 | 0.066 | 0.064 | 0.063 | 0.063 | 0.05 | < 0.5 | 0.06 | 0.06293 | 0.07 | 0.066 | 0.07 | [7 empty] | 16
| 1 | 4 | 0.06 | 0.06 | < 0.1 | 0.07 | 0.06 | 15 | 0.061 | 0.064 | 0.062 | 0.063 | 0.066 | 0.05 | < 0.5 | 0.06 | 0.06017 | 0.07 | 0.068 | 0.06 | [7 empty] | 0.063143
| 1 | 5 | 0.06 | 0.06 | < 0.1 | 0.06 | 0.06 | 14 | 0.062 | 0.057 | 0.065 | 0.063 | 0.054 | 0.04 | < 0.5 | 0.06 | 0.06142 | 0.07 | 0.067 | 0.07 | [7 empty] | 166
| 1 | 6 | 0.07 | 0.06 | < 0.1 | 0.06 | 0.05 | 14 | 0.068 | 0.06 | 0.065 | 0.065 | 0.061 | 0.05 | < 0.5 | 0.06 | 0.0603 | 0.07 | 0.065 | 0.06
| Mean |  | 0.063333 | 0.06 |  | 0.063333 | 0.065 | 14.6667 | 0.063833 | 0.062 | 0.063667 | 0.063167 | 0.062667 | 0.045 |  | 0.061667 | 0.061552 | 0.07 | 0.0665 | 0.063333
| Median |  | 0.06 | 0.06 |  | 0.065 | 0.06 | 15 | 0.063 | 0.062 | 0.0635 | 0.063 | 0.064 | 0.045 |  | 0.06 | 0.061345 | 0.07 | 0.0665 | 0.06
| Std Dev. |  | 0.005164 | 0 |  | 0.008165 | 0.017607 | 0.5164 | 0.0026394 | 0.0046904 | 0.0012111 | 0.00098319 | 0.004761 | 0.0054772 |  | 0.0040825 | 0.001285 | 0 | 0.0010488 | 0.005164
| Rel.Std.Dev. |  | 8.154% | 0% |  | 12.89% | 27.09% | 3.521% | 4.135% | 7.565% | 1.902% | 1.557% | 7.597% | 12.17% |  | 6.62% | 2.088% | 0% | 1.577% | 8.154%
| PDM3 |  | 0.3011% | -4.978% |  | 0.3011% | 2.941% | 23128% | 1.093% | -1.81% | 0.829% | 0.03717% | -0.7547% | -28.73% |  | -2.338% | -2.521% | 10.86% | 5.316% | 0.3011%
| Z-Score (Absolute) |  | 0 | 1.35 | 5.39 | 0 | 0.67 | NA | 0.2 | 0.54 | 0.13 | 0.07 | 0.27 | 7.42 | 75.52 | 0.67 | 0.72 | 2.7 | 1.28 | 0
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for K in OREAS 554 (Certified Value 0.177 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.19 | 0.1817 | 0.16 | 0.189 | 0.1675 | 0.21 | 0.21 | 0.21 | 0.16 | 0.17 | 0.18 | 0.18 | 0.18 | 0.16 | 0.1758 | 0.25082 | 0.14 | 0.16 | 0.253 | 0.19 | 0.17 | 0.16 |  |  |  | 1
| 1 | 2 | 0.19 | 0.1837 | 0.16 | 0.197 | 0.161 | 0.21 | 0.21 | 0.21 | 0.16 | 0.17 | 0.18 | 0.18 | 0.18 | 0.16 | 0.1769 | 0.23805 | 0.15 | 0.16 | 0.253 | 0.2 | 0.18 | 0.16 |  |  |  | 0
| 1 | 3 | 0.19 | 0.1766 | 0.16 | 0.214 | 0.161 | 0.2 | 0.21 | 0.21 | 0.16 | 0.17 | 0.18 | 0.17 | 0.18 | 0.15 | 0.1805 | 0.22873 | 0.15 | 0.16 | 0.254 | 0.2 | 0.18 | 0.17 |  |  |  | 16
| 1 | 4 | 0.18 | 0.1707 | 0.17 | 0.206 | 0.165 | 0.21 | 0.21 | 0.2 | 0.16 | 0.17 | 0.17 | 0.17 | 0.18 | 0.16 | 0.1766 | 0.23368 | 0.15 | 0.16 | 0.251 | 0.19 | 0.18 | 0.17 |  |  |  | 0.17697
| 1 | 5 | 0.18 | 0.1678 | 0.16 | 0.186 | 0.161 | 0.21 | 0.22 | 0.2 | 0.16 | 0.17 | 0.17 | 0.17 | 0.17 | 0.16 | 0.1755 | 0.23437 | 0.14 | 0.16 | 0.253 | 0.18 | 0.17 | 0.17 |  |  |  | 167
| 1 | 6 | 0.17 | 0.1781 | 0.17 | 0.211 | 0.167 | 0.21 | 0.21 | 0.2 | 0.15 | 0.17 | 0.17 | 0.17 | 0.18 | 0.16 | 0.177 | 0.24173 | 0.15 | 0.16 | 0.254 | 0.17 | 0.17 | 0.17
| Mean |  | 0.18333 | 0.17643 | 0.16333 | 0.2005 | 0.16375 | 0.20833 | 0.21167 | 0.205 | 0.15833 | 0.17 | 0.175 | 0.17333 | 0.17833 | 0.15833 | 0.17705 | 0.2379 | 0.14667 | 0.16 | 0.253 | 0.18833 | 0.175 | 0.16667
| Median |  | 0.185 | 0.17735 | 0.16 | 0.2015 | 0.163 | 0.21 | 0.21 | 0.205 | 0.16 | 0.17 | 0.175 | 0.17 | 0.18 | 0.16 | 0.17675 | 0.23621 | 0.15 | 0.16 | 0.253 | 0.19 | 0.175 | 0.17
| Std Dev. |  | 0.008165 | 0.0061766 | 0.005164 | 0.01164 | 0.0031265 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0040825 | 0 | 0.0054772 | 0.005164 | 0.0040825 | 0.0040825 | 0.0017942 | 0.007691 | 0.005164 | 0 | 0.0010954 | 0.01169 | 0.0054772 | 0.005164
| Rel.Std.Dev. |  | 4.454% | 3.501% | 3.162% | 5.806% | 1.909% | 1.96% | 1.929% | 2.672% | 2.578% | 0% | 3.13% | 2.979% | 2.289% | 2.578% | 1.013% | 3.233% | 3.521% | 0% | 0.433% | 6.207% | 3.13% | 3.098%
| PDM3 |  | 3.596% | -0.3033% | -7.706% | 13.3% | -7.47% | 17.72% | 19.61% | 15.84% | -10.53% | -3.939% | -1.113% | -2.055% | 0.7704% | -10.53% | 0.04521% | 34.43% | -17.12% | -9.589% | 42.96% | 6.421% | -1.113% | -5.822%
| Z-Score (Absolute) |  | 0.41 | 0.04 | 0.67 | 1.34 | 0.65 | 1.76 | 1.94 | 1.58 | 0.94 | 0.31 | 0.04 | 0.13 | 0.14 | 0.94 | 0.07 | 3.35 | 1.57 | 0.85 | 4.17 | 0.68 | 0.04 | 0.49
| Analytical results for La in OREAS 554 (Certified Value 14.2 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 23 | 24 | 28 | 29 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g |  |  |  | 1
| 1 | 1 | 13.6 | 13.319 | 8.54 | 17.12 | 13.5 | 19.1 | 15.5 | 15.2 | 14 | 14.5 | 15.7 | 15 | 14.4 | 13.7 | 7.04 | 13.3949 | 10.4 | 11.1 | 18.1 | 13.6 | 16.5 | 14.1 |  |  |  | 1
| 1 | 2 | 13.9 | 13.933 | 8.47 | 17.72 | 12.5 | 18.9 | 15.8 | 14.5 | 14 | 14.9 | 15.6 | 14.8 | 15 | 13.2 | 7.05 | 13.1176 | 10.8 | 11.2 | 18.4 | 14.1 | 16.5 | 14 |  |  |  | 0
| 1 | 3 | 13.6 | 13.191 | 8.2 | 17.23 | 12.5 | 18.5 | 15.4 | 15 | 14 | 15.2 | 15.6 | 14.5 | 15.2 | 13.3 | 7.7 | 12.9047 | 10.7 | 11.9 | 18.2 | 14 | 15.9 | 14 |  |  |  | 16
| 1 | 4 | 12.9 | 12.559 | 8.25 | 17.59 | 12.5 | 19.2 | 15.6 | 15.1 | 14 | 14.4 | 15.4 | 13.7 | 14.9 | 12.7 | 7.29 | 13.7119 | 11 | 11.9 | 18.2 | 13.8 | 16.5 | 13.6 |  |  |  | 14.2363
| 1 | 5 | 12.4 | 12.873 | 8.15 | 18.06 | 11.5 | 19.2 | 15.6 | 14.5 | 14 | 14.4 | 14.5 | 14 | 14.2 | 13.3 | 7.04 | 12.1949 | 9.9 | 12 | 18.2 | 12.6 | 16.1 | 13.8 |  |  |  | 168
| 1 | 6 | 12.2 | 13.227 | 8.23 | 17.73 | 13 | 18.8 | 15.7 | 14.6 | 13 | 14.8 | 15.1 | 14.2 | 14.9 | 13 | 7.12 | 13.9647 | 10.2 | 11.5 | 18.1 | 12.3 | 15.1 | 13.7
| Mean |  | 13.1 | 13.1837 | 8.3067 | 17.575 | 12.5833 | 18.95 | 15.6 | 14.8167 | 13.8333 | 14.7 | 15.3167 | 14.3667 | 14.7667 | 13.2 | 7.2067 | 13.2148 | 10.5 | 11.6 | 18.2 | 13.4 | 16.1 | 13.8667
| Median |  | 13.25 | 13.209 | 8.24 | 17.655 | 12.5 | 19 | 15.6 | 14.8 | 14 | 14.65 | 15.5 | 14.35 | 14.9 | 13.25 | 7.085 | 13.2563 | 10.55 | 11.7 | 18.2 | 13.7 | 16.3 | 13.9
| Std Dev. |  | 0.70427 | 0.46227 | 0.15883 | 0.3483 | 0.66458 | 0.27386 | 0.14142 | 0.31885 | 0.40825 | 0.32249 | 0.4535 | 0.49261 | 0.38297 | 0.33466 | 0.25997 | 0.6305 | 0.40988 | 0.38987 | 0.10954 | 0.76158 | 0.55136 | 0.19664
| Rel.Std.Dev. |  | 5.376% | 3.506% | 1.912% | 1.982% | 5.281% | 1.445% | 0.9065% | 2.152% | 2.951% | 2.194% | 2.961% | 3.429% | 2.593% | 2.535% | 3.607% | 4.771% | 3.904% | 3.361% | 0.6019% | 5.683% | 3.425% | 1.418%
| PDM3 |  | -7.982% | -7.394% | -41.65% | 23.45% | -11.61% | 33.11% | 9.579% | 4.077% | -2.831% | 3.257% | 7.589% | 0.9157% | 3.725% | -7.279% | -49.38% | -7.175% | -26.24% | -18.52% | 27.84% | -5.874% | 13.09% | -2.596%
| Z-Score (Absolute) |  | 0.4 | 0.36 | 2.92 | 1.95 | 0.67 | 2.67 | 0.91 | 0.5 | NA | 0.44 | 0.76 | 0.26 | 0.47 | 0.35 | 3.5 | 0.34 | 1.77 | 1.19 | 2.28 | 0.25 | 1.17 | 0
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 554 (Certified Value 18.4 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 23 | 28 | 29 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 1
|  |  | 0.25g | 25g | 01g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | [6 empty] | 1
| 1 | 1 | 19 | 18.3 | 20 | 19.6 | 21 | 20.5 | 23.5 | 20 | 16.1 | 17.3 | 19.2 | 17.6 | 16.8 | 9.1 | 12.7 | 15 | 20 | 20.5 | 16 | [6 empty] | 1
| 1 | 2 | 19 | 18.5 | 20 | 18.6 | 21 | 19.4 | 23.5 | 21 | 16.2 | 16.6 | 19.1 | 17.2 | 16.7 | 8.7 | 13.2 | 16 | 19.9 | 20.1 | 16 | [6 empty] | 0
| 1 | 3 | 19 | 17.7 | 20 | 19.8 | 23 | 19.1 | 24.3 | 21 | 16 | 17.1 | 18.5 | 17 | 16.5 | 10 | 12.7 | 16 | 19.7 | 20.3 | 17 | [6 empty] | 16
| 1 | 4 | 18 | 17.1 | 20 | 18.9 | 20 | 20.2 | 23.3 | 21 | 15.3 | 17.1 | 18.2 | 17.1 | 16.5 | 9.1 | 13.6 | 16 | 19.4 | 20.8 | 17 | [6 empty] | 18.3578
| 1 | 5 | 18 | 17.1 | 20 | 19.2 | 21 | 20.2 | 23 | 21 | 15.1 | 16.2 | 18.1 | 16.5 | 16.4 | 8.5 | 12.3 | 16 | 19.6 | 20.2 | 17 | [6 empty] | 169
| 1 | 6 | 18 | 18.3 | 20 | 19.1 | 23 | 19.6 | 23.7 | 21 | 15.5 | 16.1 | 18.4 | 17.1 | 17 | 8.8 | 12.8 | 16 | 19.6 | 19.3 | 17
| Mean |  | 18.5 | 17.8333 | 20 | 19.2 | 21.5 | 19.8333 | 23.55 | 20.8333 | 15.7 | 16.7333 | 18.5833 | 17.0833 | 16.65 | 9.0333 | 12.8833 | 15.8333 | 19.7 | 20.2 | 16.6667
| Median |  | 18.5 | 18 | 20 | 19.15 | 21 | 19.9 | 23.5 | 21 | 15.75 | 16.85 | 18.45 | 17.1 | 16.6 | 8.95 | 12.75 | 16 | 19.65 | 20.25 | 17
| Std Dev. |  | 0.54772 | 0.62823 | 0 | 0.44272 | 1.2247 | 0.5465 | 0.43704 | 0.40825 | 0.46043 | 0.50859 | 0.46224 | 0.35449 | 0.22583 | 0.52789 | 0.4535 | 0.40825 | 0.21909 | 0.50596 | 0.5164
| Rel.Std.Dev. |  | 2.961% | 3.523% | 0% | 2.306% | 5.696% | 2.755% | 1.856% | 1.96% | 2.933% | 3.039% | 2.487% | 2.075% | 1.356% | 5.844% | 3.52% | 2.578% | 1.112% | 2.505% | 3.098%
| PDM3 |  | 0.7744% | -2.857% | 8.945% | 4.587% | 17.12% | 8.037% | 28.28% | 13.48% | -14.48% | -8.849% | 1.228% | -6.943% | -9.303% | -50.79% | -29.82% | -13.75% | 7.311% | 10.03% | -9.212%
| Z-Score (Absolute) |  | 0.13 | 0.13 | 0.71 | NA | 1.28 | 0.64 | 2.07 | 1.03 | 0.95 | 0.55 | 0.16 | 0.42 | 0.58 | 3.52 | 2.04 | 0.9 | 0.59 | 0.78 | 0.58
Results from laboratory 4 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 554 (Certified Value 0.079 ppm) | [26 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 10: Lab; 11: Lab | [6 empty] | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 11; 10: 21; 11: 29 | [6 empty] | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*OES/MS; 10: AR*OES/MS; 11: AR*OES/MS | [6 empty] | 3
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.25g; 10: 0.25g; 11: 0.25g | [6 empty] | 3
| 1 | 1 | 01: 0.079; 02: 0.075; 04: < 0.1; 05: < 0.1; 06: 0.1; 10: 0.07; 11: 0.07 | [6 empty] | 1
| 1 | 2 | 01: 0.079; 02: 0.07; 04: < 0.1; 05: < 0.1; 06: 0.1; 10: 0.07; 11: 0.08 | [6 empty] | 0
| 1 | 3 | 01: 0.079; 02: 0.075; 04: < 0.1; 05: < 0.1; 06: 0.1; 10: 0.07; 11: 0.08 | [6 empty] | 16
| 1 | 4 | 01: 0.075; 02: 0.075; 04: < 0.1; 05: < 0.1; 06: 0.1; 10: 0.07; 11: 0.07 | [6 empty] | 0.079333
| 1 | 5 | 01: 0.073; 02: 0.075; 04: < 0.1; 05: < 0.1; 06: 0.1; 10: 0.07; 11: 0.07 | [6 empty] | 170
| 1 | 6 | 01: 0.075; 02: 0.075; 04: < 0.1; 05: < 0.1; 06: 0.1; 10: 0.07; 11: 0.08
| Mean |  | 01: 0.076667; 02: 0.074167; 06: 0.1; 10: 0.07; 11: 0.075
| Median |  | 01: 0.077; 02: 0.075; 06: 0.1; 10: 0.07; 11: 0.075
| Std Dev. |  | 01: 0.0026583; 02: 0.0020412; 06: 1.5202e-17; 10: 0; 11: 0.0054772
| Rel.Std.Dev. |  | 01: 3.467%; 02: 2.752%; 06: 1.52e-14%; 10: 0%; 11: 7.303%
| PDM3 |  | 01: -3.361%; 02: -6.513%; 06: 26.05%; 10: -11.76%; 11: -5.462%
| Z-Score (Absolute) |  | 01: 0.4; 02: 0; 04: 3.91; 05: 3.91; 06: 4.18; 10: 0.67; 11: 0.13
| Analytical results for Mg in OREAS 554 (Certified Value 2.9 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 27 | 28 | 29 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*AAS | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 2.9 | 2.68 | 3.01 | 2.78 | 3.0415 | 2.84 | 3.08 | 3.04 | 2.76 | 2.91 | 2.95 | 2.95 | 2.99 | 2.88 | 2.2867 | 3.0286 | 2.91 | 3 | 2.81 | 3.0697 | 2.9 | 2.76 |  |  |  | 1
| 1 | 2 | 2.94 | 2.73 | 2.96 | 2.81 | 2.983 | 2.79 | 2.9 | 2.98 | 2.74 | 2.91 | 2.9 | 2.95 | 2.97 | 2.85 | 2.2789 | 2.947 | 2.9 | 3.03 | 2.86 | 3.1592 | 2.98 | 2.76 |  |  |  | 0
| 1 | 3 | 2.92 | 2.69 | 2.91 | 2.88 | 3.0185 | 2.84 | 2.91 | 3.11 | 2.74 | 2.94 | 2.94 | 2.92 | 2.97 | 2.84 | 2.3974 | 2.6214 | 2.96 | 3.02 | 2.87 | 3.1244 | 2.89 | 2.69 |  |  |  | 16
| 1 | 4 | 2.76 | 2.6 | 2.96 | 2.85 | 3.017 | 2.83 | 2.95 | 2.98 | 2.74 | 2.89 | 2.86 | 2.87 | 2.91 | 2.83 | 2.3062 | 2.7923 | 2.94 | 2.96 | 2.83 | 3.1093 | 2.93 | 2.71 |  |  |  | 2.9015
| 1 | 5 | 2.83 | 2.57 | 2.98 | 2.88 | 2.9785 | 2.86 | 2.91 | 3 | 2.76 | 2.84 | 2.82 | 2.87 | 2.86 | 2.84 | 2.3076 | 2.7353 | 3 | 3.01 | 2.77 | 3.118 | 2.92 | 2.73 |  |  |  | 171
| 1 | 6 | 2.87 | 2.69 | 2.93 | 2.81 | 3.0315 | 2.86 | 2.82 | 2.98 | 2.78 | 2.92 | 2.85 | 2.9 | 2.97 | 2.89 | 2.3457 | 2.8874 | 2.98 | 2.98 | 2.75 | 2.8362 | 2.84 | 2.73
| Mean |  | 2.87 | 2.66 | 2.9583 | 2.835 | 3.0117 | 2.8367 | 2.9283 | 3.015 | 2.7533 | 2.9017 | 2.8867 | 2.91 | 2.945 | 2.855 | 2.3204 | 2.8353 | 2.9483 | 3 | 2.815 | 3.0695 | 2.91 | 2.73
| Median |  | 2.885 | 2.685 | 2.96 | 2.83 | 3.0177 | 2.84 | 2.91 | 2.99 | 2.75 | 2.91 | 2.88 | 2.91 | 2.97 | 2.845 | 2.3069 | 2.8399 | 2.95 | 3.005 | 2.82 | 3.1136 | 2.91 | 2.73
| Std Dev. |  | 0.066332 | 0.061319 | 0.035449 | 0.041352 | 0.025614 | 0.02582 | 0.085654 | 0.052058 | 0.01633 | 0.034303 | 0.052026 | 0.036332 | 0.049699 | 0.02429 | 0.044255 | 0.14842 | 0.0392 | 0.026077 | 0.048062 | 0.11781 | 0.046476 | 0.027568
| Rel.Std.Dev. |  | 2.311% | 2.305% | 1.198% | 1.459% | 0.8505% | 0.9102% | 2.925% | 1.727% | 0.5931% | 1.182% | 1.802% | 1.249% | 1.688% | 0.8508% | 1.907% | 5.235% | 1.33% | 0.8692% | 1.707% | 3.838% | 1.597% | 1.01%
| PDM3 |  | -1.084% | -8.322% | 1.96% | -2.29% | 3.799% | -2.233% | 0.9264% | 3.913% | -5.105% | 0.007328% | -0.5097% | 0.2945% | 1.501% | -1.601% | -20.03% | -2.279% | 1.616% | 3.396% | -2.98% | 5.79% | 0.2945% | -5.909%
| Z-Score (Absolute) |  | 0.28 | 2.68 | 0.73 | 0.68 | 1.34 | 0.66 | 0.39 | 1.38 | 1.61 | 0.09 | 0.09 | 0.18 | 0.58 | 0.45 | 6.56 | 0.67 | 0.62 | 1.21 | 0.9 | 2 | 0.18 | 1.88
| Analytical results for Mn in OREAS 554 (Certified Value 0.106 wt.%) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 27 | 28 | 29 |  |  |  | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*AAS | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.25g |  |  |  | 3
| 1 | 1 | 0.1086 | 0.1124 | 0.114 | 0.1058 | 0.1115 | 0.1097 | 0.107 | 0.112 | 0.0944 | 0.106 | 0.108 | 0.107 | 0.109 | 0.106 | 0.0946 | 0.1037 | 0.1133 | 0.103 | 0.1002 | 0.10897 | 0.1037 | 0.1022 |  |  |  | 1
| 1 | 2 | 0.1074 | 0.1151 | 0.113 | 0.1059 | 0.109 | 0.1068 | 0.106 | 0.109 | 0.0943 | 0.106 | 0.106 | 0.106 | 0.106 | 0.104 | 0.0946 | 0.099705 | 0.1132 | 0.105 | 0.1013 | 0.11084 | 0.1069 | 0.1033 |  |  |  | 0
| 1 | 3 | 0.1087 | 0.1144 | 0.11 | 0.1069 | 0.111 | 0.1092 | 0.105 | 0.115 | 0.0931 | 0.107 | 0.106 | 0.106 | 0.106 | 0.104 | 0.0962 | 0.09844 | 0.1155 | 0.103 | 0.101 | 0.10789 | 0.104 | 0.1003 |  |  |  | 16
| 1 | 4 | 0.1012 | 0.1106 | 0.112 | 0.1076 | 0.111 | 0.1092 | 0.105 | 0.11 | 0.094 | 0.104 | 0.104 | 0.105 | 0.104 | 0.105 | 0.0949 | 0.09911 | 0.112 | 0.103 | 0.1002 | 0.1085 | 0.1044 | 0.1032 |  |  |  | 0.10631
| 1 | 5 | 0.1057 | 0.111 | 0.114 | 0.1078 | 0.1095 | 0.11 | 0.107 | 0.11 | 0.0938 | 0.102 | 0.103 | 0.107 | 0.103 | 0.106 | 0.0976 | 0.098455 | 0.1161 | 0.103 | 0.1001 | 0.10813 | 0.1035 | 0.1034 |  |  |  | 172
| 1 | 6 | 0.1072 | 0.1129 | 0.116 | 0.1063 | 0.112 | 0.11 | 0.106 | 0.111 | 0.094 | 0.104 | 0.104 | 0.105 | 0.11 | 0.106 | 0.0951 | 0.099195 | 0.1159 | 0.103 | 0.1003 | 0.089804 | 0.102 | 0.1008
| Mean |  | 0.10647 | 0.11273 | 0.11317 | 0.10672 | 0.11067 | 0.10915 | 0.106 | 0.11117 | 0.093933 | 0.10483 | 0.10517 | 0.106 | 0.10633 | 0.10517 | 0.0955 | 0.099767 | 0.11433 | 0.10333 | 0.10052 | 0.10569 | 0.10408 | 0.1022
| Median |  | 0.1073 | 0.11265 | 0.1135 | 0.1066 | 0.111 | 0.10945 | 0.106 | 0.1105 | 0.094 | 0.105 | 0.105 | 0.106 | 0.106 | 0.1055 | 0.095 | 0.099153 | 0.1144 | 0.103 | 0.10025 | 0.10832 | 0.10385 | 0.1027
| Std Dev. |  | 0.0028026 | 0.0017929 | 0.0020412 | 0.00085654 | 0.001169 | 0.0012062 | 0.00089443 | 0.002137 | 0.00046332 | 0.0018348 | 0.0018348 | 0.00089443 | 0.0027325 | 0.00098319 | 0.0011866 | 0.0019856 | 0.0017166 | 0.0008165 | 0.00050365 | 0.0078534 | 0.0016043 | 0.0013579
| Rel.Std.Dev. |  | 2.632% | 1.59% | 1.804% | 0.8026% | 1.056% | 1.105% | 0.8438% | 1.922% | 0.4932% | 1.75% | 1.745% | 0.8438% | 2.57% | 0.9349% | 1.243% | 1.99% | 1.501% | 0.7902% | 0.5011% | 7.431% | 1.541% | 1.329%
| PDM3 |  | 0.1507% | 6.046% | 6.453% | 0.3859% | 4.102% | 2.675% | -0.2882% | 4.572% | -11.64% | -1.386% | -1.072% | -0.2882% | 0.02531% | -1.072% | -10.17% | -6.151% | 7.551% | -2.797% | -5.446% | -0.5798% | -2.091% | -3.863%
| Z-Score (Absolute) |  | 0.14 | 1.6 | 1.7 | 0.2 | 1.12 | 0.77 | 0.04 | 1.23 | 2.76 | 0.23 | 0.16 | 0.04 | 0.11 | 0.16 | 2.4 | 1.41 | 1.97 | 0.58 | 1.24 | 0.04 | 0.41 | 0.85
| Analytical results for Mo in OREAS 554 (Certified Value 3.32 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 3.41 | 3.4 | 4 | 3 | 4 | 3.2 | 3.44 | 3.49 | 3 | 3.4 | 3.39 | 3.38 | 3.5 | 3.46 | 3.1 | 2.9406 | 3.38 | 3.2769 | 3.2 | 3.29 | 3.5 | 3.39 | 2.75 |  |  | 1
| 1 | 2 | 3.41 | 3.5 | 4 | 3 | 4 | 3.2 | 3.35 | 3.49 | 3 | 3.45 | 3.32 | 3.48 | 3.66 | 3.25 | 3.1 | 3.168 | 3.22 | 3.2339 | 3.2 | 3.18 | 3.53 | 3.47 | 2.52 |  |  | 0
| 1 | 3 | 3.12 | 3.3 | 3.8 | 3.1 | 4 | 3.2 | 3.3 | 3.51 | 3 | 3.39 | 3.29 | 3.37 | 3.58 | 3.15 | 3.6 | 2.7645 | 3.38 | 3.1456 | 3.1 | 3.13 | 3.54 | 3.52 | 2.79 |  |  | 16
| 1 | 4 | 3.06 | 3.4 | 3.8 | 3 | 4 | 3.2 | 3.37 | 3.46 | 2 | 3.37 | 3.53 | 3.42 | 3.6 | 3.46 | 3.4 | 3.0441 | 3.44 | 3.3035 | 3.1 | 3.35 | 3.52 | 3.35 | 2.74 |  |  | 3.3198
| 1 | 5 | 3.3 | 3.3 | 3.8 | 3.1 | 4 | 3.3 | 3.42 | 3.5 | 2 | 3.37 | 3.36 | 3.37 | 3.56 | 3.36 | 3.1 | 3.1754 | 3.36 | 3.4438 | 3.1 | 3.03 | 3.6 | 3.42 | 2.5 |  |  | 173
| 1 | 6 | 3.1 | 3.5 | 4 | 3 | 4 | 3.1 | 3.27 | 3.47 | 3 | 3.46 | 3.42 | 3.38 | 3.71 | 3.48 | 3.1 | 3.0596 | 3.64 | 3.3479 | 3.2 | 3.22 | 3.43 | 3.58 | 2.64
| Mean |  | 3.2333 | 3.4 | 3.9 | 3.0333 | 4 | 3.2 | 3.3583 | 3.4867 | 2.6667 | 3.4067 | 3.385 | 3.4 | 3.6017 | 3.36 | 3.2333 | 3.0254 | 3.4033 | 3.2919 | 3.15 | 3.2 | 3.52 | 3.455 | 2.6567
| Median |  | 3.21 | 3.4 | 3.9 | 3 | 4 | 3.2 | 3.36 | 3.49 | 3 | 3.395 | 3.375 | 3.38 | 3.59 | 3.41 | 3.1 | 3.0519 | 3.38 | 3.2902 | 3.15 | 3.2 | 3.525 | 3.445 | 2.69
| Std Dev. |  | 0.15971 | 0.089443 | 0.10954 | 0.05164 | 0 | 0.063246 | 0.066156 | 0.018619 | 0.5164 | 0.039328 | 0.085029 | 0.043359 | 0.074409 | 0.13461 | 0.21602 | 0.15466 | 0.13706 | 0.10131 | 0.054772 | 0.11419 | 0.055498 | 0.085499 | 0.12404
| Rel.Std.Dev. |  | 4.939% | 2.631% | 2.809% | 1.702% | 0% | 1.976% | 1.97% | 0.534% | 19.36% | 1.154% | 2.512% | 1.275% | 2.066% | 4.006% | 6.681% | 5.112% | 4.027% | 3.078% | 1.739% | 3.569% | 1.577% | 2.475% | 4.669%
| PDM3 |  | -2.603% | 2.417% | 17.48% | -8.628% | 20.49% | -3.608% | 1.162% | 5.028% | -19.67% | 2.618% | 1.965% | 2.417% | 8.492% | 1.212% | -2.603% | -8.868% | 2.517% | -0.8383% | -5.114% | -3.608% | 6.032% | 4.074% | -19.97%
| Z-Score (Absolute) |  | 0.67 | 0.21 | 2.87 | 1.74 | NA | 0.85 | 0.01 | 0.67 | NA | 0.25 | 0.13 | 0.21 | 1.29 | 0 | 0.67 | 1.78 | 0.23 | 0.36 | 1.12 | 0.85 | 0.85 | 0.51 | 3.74
Results from laboratories 6 & 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 554 (Indicative Value 0.018 wt.%) | [26 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | [6 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 3
|  |  | 0.25g | 25g | 01g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.02 | 0.02 | 0.03 | 0.017 | < 0.01 | 0.015 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.02 | NA | 0.026149 | < 0.01 | 0.01 | 0.008 | 0.04 | < 0.01 | [6 empty] | 1
| 1 | 2 | 0.02 | < 0.01 | 0.03 | 0.015 | < 0.01 | 0.017 | < 0.01 | < 0.01 | 0.01 | 0.03 | 0.01 | 0.02 | NA | 0.023907 | < 0.01 | 0.01 | 0.008 | 0.03 | < 0.01 | [6 empty] | 0
| 1 | 3 | 0.02 | 0.02 | 0.03 | 0.0145 | < 0.01 | 0.017 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.02 | NA | 0.021385 | < 0.01 | 0.01 | 0.008 | 0.03 | < 0.01 | [6 empty] | 16
| 1 | 4 | 0.02 | < 0.01 | 0.03 | 0.0155 | < 0.01 | 0.017 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.01 | NA | 0.020872 | < 0.01 | 0.01 | 0.008 | 0.04 | < 0.01 | [6 empty] | 0.017971
| 1 | 5 | 0.02 | < 0.01 | 0.03 | 0.0155 | < 0.01 | 0.019 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.01 | NA | 0.022828 | < 0.01 | 0.01 | 0.007 | 0.03 | < 0.01 | [6 empty] | 29
| 1 | 6 | 0.02 | < 0.01 | 0.03 | 0.0165 | < 0.01 | 0.017 | < 0.01 | 0.01 | 0.01 | 0.03 | 0.01 | 0.01 | NA | 0.022399 | < 0.01 | 0.01 | 0.007 | 0.03 | < 0.01
| Mean |  | 0.02 | 0.02 | 0.03 | 0.015667 |  | 0.017 |  | 0.01 | 0.01 | 0.03 | 0.01 | 0.015 |  | 0.022923 |  | 0.01 | 0.0076667 | 0.033333
| Median |  | 0.02 | 0.02 | 0.03 | 0.0155 |  | 0.017 |  | 0.01 | 0.01 | 0.03 | 0.01 | 0.015 |  | 0.022614 |  | 0.01 | 0.008 | 0.03
| Std Dev. |  | 0 | 0 | 0 | 0.00093095 |  | 0.0012649 |  | 0 | 0 | 0 | 0 | 0.0054772 |  | 0.0019098 |  | 0 | 0.0005164 | 0.005164
| Rel.Std.Dev. |  | 0% | 0% | 0% | 5.942% |  | 7.441% |  | 0% | 0% | 0% | 0% | 36.51% |  | 8.331% |  | 0% | 6.736% | 15.49%
| PDM3 |  | 11.29% | 11.29% | 66.94% | -12.82% |  | -5.402% |  | -44.35% | -44.35% | 66.94% | -44.35% | -16.53% |  | 27.56% |  | -44.35% | -57.34% | 85.49%
| Z-Score (Absolute) |  | 1.35 | 0 | 2.7 | 0.76 | 0.67 | 0.94 | 0.67 | 0.11 | 0 | 2.7 | 0 | 0.67 | NA | 1.74 | 0.67 | 0 | 0.31 | 3.15 | 0.67
| Analytical results for Nb in OREAS 554 (Indicative Value < 0.05 ppm) | [26 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.03 | < 0.2 | < 1 | < 0.1 | 0.07 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | NA | < 0.01 | < 0.05 | < 0.2 | < 0.05 | < 0.05 | 0.03 | [7 empty] | 1
| 1 | 2 | 0.04 | < 0.2 | < 1 | < 0.1 | 0.06 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | NA | 0.019963 | < 0.05 | < 0.2 | < 0.05 | < 0.05 | 0.03 | [7 empty] | 0
| 1 | 3 | 0.03 | < 0.2 | < 1 | < 0.1 | 0.1 | < 5 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | NA | < 0.01 | < 0.05 | < 0.2 | < 0.05 | < 0.05 | 0.02 | [7 empty] | 16
| 1 | 4 | 0.03 | < 0.2 | < 1 | < 0.1 | < 0.05 | < 5 | < 0.05 | 0.1 | < 0.05 | < 0.05 | < 0.05 | NA | < 0.01 | < 0.05 | < 0.2 | < 0.05 | < 0.05 | 0.03 | [7 empty] | < 0.05
| 1 | 5 | 0.03 | < 0.2 | < 1 | < 0.1 | < 0.05 | < 5 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.05 | NA | 0.028015 | < 0.05 | < 0.2 | < 0.05 | < 0.05 | 0.02 | [7 empty] | 30
| 1 | 6 | 0.04 | < 0.2 | < 1 | < 0.1 | < 0.05 | < 5 | < 0.05 | < 0.05 | < 0.05 | 0.08 | < 0.05 | NA | < 0.01 | < 0.05 | < 0.2 | 0.05 | < 0.05 | 0.28
| Mean |  | 02: 0.033333; 11: 0.076667; 14: 0.075; 16: 0.08; 19: 0.023989; 28: 0.05; 30: 0.068333
| Median |  | 02: 0.03; 11: 0.07; 14: 0.075; 16: 0.08; 19: 0.023989; 28: 0.05; 30: 0.03
| Std Dev. |  | 02: 0.005164; 11: 0.020817; 14: 0.035355; 19: 0.0056942; 30: 0.10381
| Rel.Std.Dev. |  | 02: 15.49%; 11: 27.15%; 14: 47.14%; 19: 23.74%; 30: 151.9%
PDM3
| Z-Score (Absolute) |  | 0.06 | 4.84 | 34.27 | 1.16 | 1.23 | 181.39 | 0.67 | 0.55 | 0.67 | 0 | 0.67 | NA | 1.68 | 0.67 | 4.84 | 0.37 | 0.67 | 2.51
| Analytical results for Nd in OREAS 554 (Certified Value 13 ppm) | [26 empty] | Certified
Nd | Round | Replicate | 02: Lab; 05: Lab; 06: Lab; 10: Lab; 11: Lab; 12: Lab | [7 empty] | 1
| No | No | 02: 02; 05: 04; 06: 06; 10: 10; 11: 19; 12: 30 | [7 empty] | ppm
|  |  | 02: AR*MS; 05: AR*MS; 06: AR*MS; 10: AR*OES/MS; 11: AR*OES/MS; 12: AR*MS | [7 empty] | 1
|  |  | 02: 25g; 05: 01-05g; 06: 25g; 10: 0.5g; 11: 01g; 12: 0.5g | [7 empty] | 1
| 1 | 1 | 02: 12.915; 05: 10.2; 06: 13.5; 10: 14.3; 11: 14.5509; 12: 6.25 | [7 empty] | 1
| 1 | 2 | 02: 13.568; 05: 11; 06: 12.5; 10: 14.9; 11: 14.1049; 12: 6.07 | [7 empty] | 0
| 1 | 3 | 02: 12.828; 05: 10.2; 06: 12.5; 10: 14.7; 11: 14.1088; 12: 6.12 | [7 empty] | 16
| 1 | 4 | 02: 12.299; 05: 10; 06: 12.5; 10: 14.9; 11: 15.1674; 12: 6.96 | [7 empty] | 12.9641
| 1 | 5 | 02: 12.639; 05: 9.79; 06: 11.5; 10: 15; 11: 13.1004; 12: 6.05 | [7 empty] | 174
| 1 | 6 | 02: 12.88; 05: 10.2; 06: 13; 10: 14.9; 11: 15.1716; 12: 6.22
| Mean |  | 02: 12.8548; 05: 10.2317; 06: 12.5833; 10: 14.7833; 11: 14.3673; 12: 6.2783
| Median |  | 02: 12.854; 05: 10.2; 06: 12.5; 10: 14.9; 11: 14.3298; 12: 6.17
| Std Dev. |  | 02: 0.41692; 05: 0.41039; 06: 0.66458; 10: 0.25626; 11: 0.7825; 12: 0.34336
| Rel.Std.Dev. |  | 02: 3.243%; 05: 4.011%; 06: 5.281%; 10: 1.733%; 11: 5.446%; 12: 5.469%
| PDM3 |  | 02: -0.8428%; 05: -21.08%; 06: -2.937%; 10: 14.03%; 11: 10.82%; 12: -51.57%
| Z-Score (Absolute) |  | 02: 0.05; 05: 0.9; 06: 0.05; 10: 0.75; 11: 0.6; 12: 2.34
| Analytical results for Ni in OREAS 554 (Certified Value 28.4 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 30 | 27.6 | 31 | 28 | 27 | 30.7 | 27.7 | 27.8 | 25 | 27 | 30.6 | 25.9 | 29.9 | 28.7 | NA | 28.6457 | 25.3 | 28.7695 | 28 | 26.5 | 30.1 | 31.7 | 30.4 |  |  | 1
| 1 | 2 | 29 | 28.9 | 33 | 29 | 29 | 31.7 | 26.7 | 27.2 | 25 | 27.1 | 30.8 | 26.2 | 30.1 | 28.1 | NA | 28.6288 | 25.7 | 30.1461 | 27 | 23.6 | 30.5 | 31.7 | 30.2 |  |  | 0
| 1 | 3 | 30 | 28.2 | 30 | 29 | 29 | 29.3 | 26.2 | 27.1 | 25 | 27.2 | 29.6 | 26.2 | 29.9 | 27.9 | NA | 24.8575 | 31.4 | 29.8625 | 27 | 24 | 29.7 | 32.2 | 31.2 |  |  | 16
| 1 | 4 | 28 | 27.6 | 32 | 29 | 28 | 28.9 | 27.3 | 26.1 | 25 | 26 | 31.1 | 25.5 | 28.6 | 28.9 | NA | 26.4662 | 28 | 28.6673 | 27 | 24.2 | 30.1 | 31.7 | 31.5 |  |  | 28.388
| 1 | 5 | 33 | 27.3 | 31 | 30 | 27 | 29.2 | 27.5 | 26.8 | 25 | 26.4 | 30.3 | 26.3 | 26.5 | 27 | NA | 25.2541 | 26.8 | 29.5171 | 28 | 22.8 | 29.5 | 32.3 | 30.7 |  |  | 175
| 1 | 6 | 30 | 28 | 32 | 28 | 28 | 31.2 | 27 | 27.6 | 25 | 26.5 | 31.6 | 25.9 | 30.1 | 27 | NA | 27.4504 | 31.7 | 29.9949 | 26 | 23.1 | 29 | 32.2 | 31.4
| Mean |  | 30 | 27.9333 | 31.5 | 28.8333 | 28 | 30.1667 | 27.0667 | 27.1 | 25 | 26.7 | 30.6667 | 26 | 29.1833 | 27.9333 |  | 26.8838 | 28.15 | 29.4929 | 27.1667 | 24.0333 | 29.8167 | 31.9667 | 30.9
| Median |  | 30 | 27.8 | 31.5 | 29 | 28 | 30 | 27.15 | 27.15 | 25 | 26.75 | 30.7 | 26.05 | 29.9 | 28 |  | 26.9583 | 27.4 | 29.6898 | 27 | 23.8 | 29.9 | 31.95 | 30.95
| Std Dev. |  | 1.6733 | 0.57155 | 1.0488 | 0.75277 | 0.89443 | 1.1827 | 0.55377 | 0.60663 | 0 | 0.47329 | 0.68605 | 0.29665 | 1.4317 | 0.81158 |  | 1.6375 | 2.7977 | 0.63574 | 0.75277 | 1.3186 | 0.53072 | 0.29439 | 0.54406
| Rel.Std.Dev. |  | 5.578% | 2.046% | 3.33% | 2.611% | 3.194% | 3.92% | 2.046% | 2.238% | 0% | 1.773% | 2.237% | 1.141% | 4.906% | 2.905% |  | 6.091% | 9.938% | 2.156% | 2.771% | 5.486% | 1.78% | 0.9209% | 1.761%
| PDM3 |  | 5.678% | -1.602% | 10.96% | 1.569% | -1.367% | 6.265% | -4.655% | -4.537% | -11.93% | -5.946% | 8.027% | -8.412% | 2.802% | -1.602% |  | -5.299% | -0.8385% | 3.892% | -4.302% | -15.34% | 5.033% | 12.61% | 8.849%
| Z-Score (Absolute) |  | 0.93 | 0.07 | 1.65 | 0.37 | 0.04 | 1.01 | 0.49 | 0.47 | 1.48 | 0.66 | 1.25 | 1 | 0.54 | 0.07 | NA | 0.58 | 0.04 | 0.68 | 0.44 | 1.95 | 0.84 | 1.88 | 1.36
| Analytical results for P in OREAS 554 (Certified Value 0.026 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | [6 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 3
|  |  | 0.25g | 25g | 01g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.03 | 0.0262 | 0.028 | 0.0285 | 0.031 | 0.027 | 0.0337 | 0.027 | 0.025 | 0.027 | 0.026 | 0.027 | 0.0253 | 0.021115 | 0.0258 | 0.0297 | 0.026 | 0.0264 | 0.0255 | [6 empty] | 1
| 1 | 2 | 0.03 | 0.0268 | 0.026 | 0.0265 | 0.029 | 0.026 | 0.0333 | 0.027 | 0.025 | 0.027 | 0.026 | 0.027 | 0.0257 | 0.025 | 0.0264 | 0.0306 | 0.026 | 0.027 | 0.026 | [6 empty] | 0
| 1 | 3 | 0.03 | 0.0258 | 0.027 | 0.0285 | 0.036 | 0.026 | 0.0338 | 0.027 | 0.025 | 0.026 | 0.027 | 0.027 | 0.0264 | 0.021198 | 0.0286 | 0.0297 | 0.026 | 0.0264 | 0.0263 | [6 empty] | 16
| 1 | 4 | 0.03 | 0.0262 | 0.028 | 0.029 | 0.035 | 0.027 | 0.0326 | 0.027 | 0.024 | 0.026 | 0.026 | 0.026 | 0.0255 | 0.023539 | 0.0262 | 0.0293 | 0.026 | 0.0266 | 0.0257 | [6 empty] | 0.026448
| 1 | 5 | 0.03 | 0.0261 | 0.029 | 0.027 | 0.034 | 0.026 | 0.033 | 0.027 | 0.024 | 0.026 | 0.026 | 0.027 | 0.0258 | 0.024095 | 0.0275 | 0.0285 | 0.025 | 0.0268 | 0.0254 | [6 empty] | 176
| 1 | 6 | 0.03 | 0.0252 | 0.027 | 0.027 | 0.032 | 0.027 | 0.0335 | 0.027 | 0.025 | 0.027 | 0.026 | 0.027 | 0.0258 | 0.026203 | 0.0261 | 0.0297 | 0.025 | 0.0255 | 0.025
| Mean |  | 0.03 | 0.02605 | 0.0275 | 0.02775 | 0.032833 | 0.0265 | 0.033317 | 0.027 | 0.024667 | 0.0265 | 0.026167 | 0.026833 | 0.02575 | 0.023525 | 0.026767 | 0.029583 | 0.025667 | 0.02645 | 0.02565
| Median |  | 0.03 | 0.02615 | 0.0275 | 0.02775 | 0.033 | 0.0265 | 0.0334 | 0.027 | 0.025 | 0.0265 | 0.026 | 0.027 | 0.02575 | 0.023817 | 0.0263 | 0.0297 | 0.026 | 0.0265 | 0.0256
| Std Dev. |  | 0 | 0.0005282 | 0.0010488 | 0.0010368 | 0.0026394 | 0.00054772 | 0.0004535 | 0 | 0.0005164 | 0.00054772 | 0.00040825 | 0.00040825 | 0.00037283 | 0.0020442 | 0.0010708 | 0.0006824 | 0.0005164 | 0.00052058 | 0.00045935
| Rel.Std.Dev. |  | 0% | 2.028% | 3.814% | 3.736% | 8.039% | 2.067% | 1.361% | 0% | 2.094% | 2.067% | 1.56% | 1.521% | 1.448% | 8.689% | 4.001% | 2.307% | 2.012% | 1.968% | 1.791%
| PDM3 |  | 13.43% | -1.507% | 3.976% | 4.921% | 24.14% | 0.1948% | 25.97% | 2.085% | -6.737% | 0.1948% | -1.066% | 1.455% | -2.641% | -11.05% | 1.203% | 11.85% | -2.956% | 0.005727% | -3.019%
| Z-Score (Absolute) |  | 2.83 | 0.36 | 0.81 | 1.01 | 5.12 | 0 | 5.52 | 0.4 | 1.48 | 0 | 0.27 | 0.27 | 0.61 | 2.41 | 0.22 | 2.49 | 0.67 | 0.04 | 0.69
| Analytical results for Pb in OREAS 554 (Certified Value 11 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 1
| 1 | 1 | 10.6 | 11.3 | 10 | 11.5 | 11 | 12 | 12.5 | 14.3 | 9 | 10.7 | 11.2 | 10 | 11 | 11.2 | 6 | 11.0795 | 9.2 | 10.6873 | 11.9 | 10.81 | 12.6 | 9.8 | 10.7 |  |  | 1
| 1 | 2 | 10.5 | 12.3 | 10 | 11.4 | 10 | 12.1 | 13.6 | 14.4 | 9 | 10.6 | 11.2 | 10.1 | 11.1 | 10.5 | 6 | 9.9948 | 8.9 | 9.9782 | 12.5 | 10.91 | 12.6 | 9.5 | 8.24 |  |  | 0
| 1 | 3 | 10.4 | 11.3 | 9 | 11.2 | 10 | 11.8 | 12.3 | 13.4 | 9 | 10.9 | 11.1 | 10.5 | 11.1 | 10.3 | 6 | 9.0237 | 9.4 | 9.9488 | 12.3 | 11.34 | 12.1 | 9.5 | 8.64 |  |  | 16
| 1 | 4 | 9.9 | 10.9 | 9 | 11.5 | 11 | 12.1 | 12.8 | 13.6 | 9 | 10.4 | 10.9 | 10 | 11.2 | 10.5 | 5 | 9.3545 | 9.6 | 9.7924 | 12 | 11.02 | 12.7 | 9.4 | 8.84 |  |  | 10.9759
| 1 | 5 | 10 | 11 | 9 | 11.1 | 10 | 12.1 | 12.7 | 14.7 | 10 | 10.5 | 10.7 | 9.8 | 10.7 | 11.2 | 5 | 9.8006 | 9.5 | 10.101 | 12.2 | 11.24 | 12.4 | 9.7 | 8.42 |  |  | 177
| 1 | 6 | 9.9 | 11.5 | 10 | 11.3 | 11 | 11.9 | 12.8 | 15 | 10 | 10.8 | 11.2 | 9.6 | 11.4 | 10.5 | 6 | 9.9266 | 9.6 | 10.8 | 12.4 | 11.34 | 11.9 | 9.7 | 8.53
| Mean |  | 10.2167 | 11.3833 | 9.5 | 11.3333 | 10.5 | 12 | 12.7833 | 14.2333 | 9.3333 | 10.65 | 11.05 | 10 | 11.0833 | 10.7 | 5.6667 | 9.8633 | 9.3667 | 10.218 | 12.2167 | 11.11 | 12.3833 | 9.6 | 8.895
| Median |  | 10.2 | 11.3 | 9.5 | 11.35 | 10.5 | 12.05 | 12.75 | 14.35 | 9 | 10.65 | 11.15 | 10 | 11.1 | 10.5 | 6 | 9.8636 | 9.45 | 10.0396 | 12.25 | 11.13 | 12.5 | 9.6 | 8.585
| Std Dev. |  | 0.31885 | 0.49967 | 0.54772 | 0.1633 | 0.54772 | 0.12649 | 0.4446 | 0.62183 | 0.5164 | 0.18708 | 0.20736 | 0.30332 | 0.23166 | 0.39497 | 0.5164 | 0.70259 | 0.27325 | 0.42041 | 0.23166 | 0.22839 | 0.31885 | 0.15492 | 0.90708
| Rel.Std.Dev. |  | 3.121% | 4.389% | 5.766% | 1.441% | 5.216% | 1.054% | 3.478% | 4.369% | 5.533% | 1.757% | 1.877% | 3.033% | 2.09% | 3.691% | 9.113% | 7.123% | 2.917% | 4.114% | 1.896% | 2.056% | 2.575% | 1.614% | 10.2%
| PDM3 |  | -6.917% | 3.712% | -13.45% | 3.257% | -4.336% | 9.331% | 16.47% | 29.68% | -14.97% | -2.969% | 0.6752% | -8.891% | 0.9789% | -2.514% | -48.37% | -10.14% | -14.66% | -6.905% | 11.3% | 1.222% | 12.82% | -12.54% | -18.96%
| Z-Score (Absolute) |  | 0.59 | 0.24 | NA | 0.2 | NA | 0.67 | 1.23 | 2.26 | NA | 0.28 | 0 | 0.75 | 0.02 | 0.25 | NA | 0.84 | 1.19 | 0.59 | 0.83 | 0.04 | 0.95 | 1.03 | 1.53
Results from laboratories 4, 6, 12 & 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 554 (Indicative Value < 10 ppb) | [26 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab; 04: Lab |  |  | 1
| No | No | 01: 02; 02: 05; 04: 06 |  |  | ppb
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS |  |  | 1
|  |  | 01: 25g; 02: 01g; 04: 25g |  |  | 1
| 1 | 1 | 01: < 10; 02: < 10; 04: < 50 |  |  | 1
| 1 | 2 | 01: < 10; 02: < 10; 04: < 50 |  |  | 0
| 1 | 3 | 01: < 10; 02: < 10; 04: < 50 |  |  | 16
| 1 | 4 | 01: < 10; 02: < 10; 04: < 50 |  |  | < 10
| 1 | 5 | 01: < 10; 02: < 10; 04: < 50 |  |  | 27
| 1 | 6 | 01: < 10; 02: < 10; 04: < 50
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 04
| Analytical results for Pr in OREAS 554 (Certified Value 2.82 ppm) | [26 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab |  |  | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 30 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS |  |  | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.5g |  |  | 2
| 1 | 1 | 01: 3.345; 02: 2.39; 04: 3.4; 05: 3.8; 06: 1.62 |  |  | 1
| 1 | 2 | 01: 3.523; 02: 2.52; 04: 3.1; 05: 3.9; 06: 1.56 |  |  | 0
| 1 | 3 | 01: 3.316; 02: 2.21; 04: 3.1; 05: 3.7; 06: 1.6 |  |  | 16
| 1 | 4 | 01: 3.158; 02: 2.22; 04: 3.1; 05: 3.8; 06: 1.57 |  |  | 2.8174
| 1 | 5 | 01: 3.287; 02: 2.21; 04: 2.9; 05: 3.8; 06: 1.52 |  |  | 178
| 1 | 6 | 01: 3.342; 02: 2.22; 04: 3.3; 05: 3.7; 06: 1.58
| Mean |  | 01: 3.3285; 02: 2.295; 04: 3.15; 05: 3.7833; 06: 1.575
| Median |  | 01: 3.329; 02: 2.22; 04: 3.1; 05: 3.8; 06: 1.575
| Std Dev. |  | 01: 0.11767; 02: 0.13065; 04: 0.17607; 05: 0.075277; 06: 0.034496
| Rel.Std.Dev. |  | 01: 3.535%; 02: 5.693%; 04: 5.589%; 05: 1.99%; 06: 2.19%
| PDM3 |  | 01: 18.14%; 02: -18.54%; 04: 11.81%; 05: 34.29%; 06: -44.1%
| Z-Score (Absolute) |  | 01: 0.19; 02: 0.91; 04: 0; 05: 0.67; 06: 1.68
| Analytical results for Pt in OREAS 554 (Indicative Value 13.3 ppb) | [26 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab; 04: Lab |  |  | 1
| No | No | 01: 02; 02: 05; 04: 06 |  |  | ppb
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS |  |  | 1
|  |  | 01: 25g; 02: 01g; 04: 25g |  |  | 1
| 1 | 1 | 01: < 5; 02: < 5; 04: 10 |  |  | 1
| 1 | 2 | 01: < 5; 02: < 5; 04: 10 |  |  | 0
| 1 | 3 | 01: < 5; 02: < 5; 04: 10 |  |  | 16
| 1 | 4 | 01: < 5; 02: < 5; 04: 20 |  |  | 13.3333
| 1 | 5 | 01: < 5; 02: < 5; 04: 10 |  |  | 28
| 1 | 6 | 01: < 5; 02: < 5; 04: 20
| Mean |  | 04: 13.3333
| Median |  | 04: 10
| Std Dev. |  | 04: 5.164
| Rel.Std.Dev. |  | 04: 38.73%
| PDM3 |  | 04: 2.442e-13%
Z-Score (Absolute) | bar: 01 to 04
| Analytical results for Rb in OREAS 554 (Certified Value 8.83 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [7 empty] | 3
| 1 | 1 | 9.4 | 9.68 | 8.6 | 10.06 | 9 | 11.3 | 9.1 | 9.3 | 9.1 | 8.8 | 8.9 | 7 | 11.1957 | 7.5 | 14.2 | 11.6 | 8.7 | 7.11 | [7 empty] | 1
| 1 | 2 | 9.5 | 10.19 | 8.4 | 10.18 | 8.5 | 11.4 | 9.1 | 8.9 | 9.6 | 9.5 | 8.7 | 7 | 12.59 | 7.7 | 13.5 | 11.3 | 8.6 | 7.01 | [7 empty] | 0
| 1 | 3 | 9.1 | 9.55 | 8.8 | 10.97 | 8.5 | 11.4 | 9.3 | 9.2 | 9.1 | 9.6 | 8.7 | 8 | 11.3218 | 8.1 | 14 | 11.6 | 8.5 | 7.15 | [7 empty] | 16
| 1 | 4 | 8.8 | 9.36 | 8.6 | 10.3 | 8.5 | 11.5 | 8.8 | 9.9 | 8.9 | 9.2 | 7.8 | 8 | 11.1891 | 8 | 14.4 | 11.2 | 8.7 | 7 | [7 empty] | 8.8308
| 1 | 5 | 8.5 | 9.28 | 8.5 | 11.25 | 8 | 11.8 | 8.7 | 9.4 | 8.8 | 8.3 | 7.8 | 7 | 11.4077 | 8 | 14.3 | 11.9 | 8.6 | 7.08 | [7 empty] | 179
| 1 | 6 | 8.3 | 9.67 | 8.6 | 11.08 | 8.5 | 11.2 | 9.2 | 9.4 | 9 | 9 | 8 | 7 | 12.3031 | 8 | 14.6 | 11.1 | 8.5 | 7.32
| Mean |  | 8.9333 | 9.6217 | 8.5833 | 10.64 | 8.5 | 11.4333 | 9.0333 | 9.35 | 9.0833 | 9.0667 | 8.3167 | 7.3333 | 11.6679 | 7.8833 | 14.1667 | 11.45 | 8.6 | 7.1117
| Median |  | 8.95 | 9.61 | 8.6 | 10.635 | 8.5 | 11.4 | 9.1 | 9.35 | 9.05 | 9.1 | 8.35 | 7 | 11.3647 | 8 | 14.25 | 11.45 | 8.6 | 7.095
| Std Dev. |  | 0.48442 | 0.32221 | 0.13292 | 0.51734 | 0.31623 | 0.20656 | 0.23381 | 0.32711 | 0.27869 | 0.48028 | 0.50365 | 0.5164 | 0.6154 | 0.23166 | 0.38297 | 0.30166 | 0.089443 | 0.1172
| Rel.Std.Dev. |  | 5.423% | 3.349% | 1.549% | 4.862% | 3.72% | 1.807% | 2.588% | 3.498% | 3.068% | 5.297% | 6.056% | 7.042% | 5.274% | 2.939% | 2.703% | 2.635% | 1.04% | 1.648%
| PDM3 |  | 1.161% | 8.956% | -2.802% | 20.49% | -3.746% | 29.47% | 2.294% | 5.88% | 2.86% | 2.671% | -5.822% | -16.96% | 32.13% | -10.73% | 60.42% | 29.66% | -2.613% | -19.47%
| Z-Score (Absolute) |  | 0.16 | 0.66 | 0.58 | 1.87 | 0.67 | 2.82 | 0.04 | 0.34 | 0.02 | 0 | 0.89 | NA | 3.1 | 1.41 | 6.07 | 2.84 | 0.56 | 2.33
Results from laboratory 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Re in OREAS 554 (Certified Value 0.045 ppm) | [26 empty] | Certified
Re | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab |  |  | 1
| No | No | 01: 02; 02: 05; 04: 06; 05: 10; 06: 13; 09: 14; 10: 15; 11: 16; 12: 17; 13: 21; 14: 23; 15: 28; 16: 29 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS; 09: AR*MS; 10: AR*MS; 11: AR*MS; 12: AR*MS; 13: AR*OES/MS; 14: AR*OES/MS; 15: AR*OES/MS; 16: AR*OES/MS |  |  | 3
|  |  | 01: 25g; 02: 01g; 04: 25g; 05: 0.5g; 06: 0.5g; 09: 0.5g; 10: 0.5g; 11: 0.5g; 12: 0.5g; 13: 0.25g; 14: 01g; 15: 0.5g; 16: 0.25g |  |  | 3
| 1 | 1 | 01: 0.04; 02: < 0.05; 04: 0.05; 05: 0.04; 06: 0.044; 09: 0.039; 10: 0.046; 11: 0.044; 12: 0.047; 13: < 0.1; 14: < 0.05; 15: 0.044; 16: 0.048 |  |  | 1
| 1 | 2 | 01: 0.043; 02: < 0.05; 04: 0.05; 05: 0.04; 06: 0.043; 09: 0.046; 10: 0.048; 11: 0.045; 12: 0.042; 13: < 0.1; 14: < 0.05; 15: 0.046; 16: 0.042 |  |  | 0
| 1 | 3 | 01: 0.043; 02: < 0.05; 04: 0.05; 05: 0.039; 06: 0.044; 09: 0.044; 10: 0.045; 11: 0.046; 12: 0.043; 13: < 0.1; 14: < 0.05; 15: 0.045; 16: 0.046 |  |  | 16
| 1 | 4 | 01: 0.047; 02: < 0.05; 04: 0.05; 05: 0.039; 06: 0.042; 09: 0.048; 10: 0.044; 11: 0.043; 12: 0.046; 13: < 0.1; 14: < 0.05; 15: 0.045; 16: 0.051 |  |  | 0.044583
| 1 | 5 | 01: 0.04; 02: < 0.05; 04: 0.05; 05: 0.041; 06: 0.043; 09: 0.049; 10: 0.047; 11: 0.043; 12: 0.049; 13: < 0.1; 14: < 0.05; 15: 0.045; 16: 0.047 |  |  | 180
| 1 | 6 | 01: 0.04; 02: < 0.05; 04: 0.05; 05: 0.041; 06: 0.043; 09: 0.043; 10: 0.044; 11: 0.043; 12: 0.047; 13: < 0.1; 14: < 0.05; 15: 0.043; 16: 0.045
| Mean |  | 01: 0.042167; 04: 0.05; 05: 0.04; 06: 0.043167; 09: 0.044833; 10: 0.045667; 11: 0.044; 12: 0.045667; 15: 0.044667; 16: 0.0465
| Median |  | 01: 0.0415; 04: 0.05; 05: 0.04; 06: 0.043; 09: 0.045; 10: 0.0455; 11: 0.0435; 12: 0.0465; 15: 0.045; 16: 0.0465
| Std Dev. |  | 01: 0.0027869; 04: 7.6012e-18; 05: 0.00089443; 06: 0.00075277; 09: 0.003656; 10: 0.001633; 11: 0.0012649; 12: 0.0026583; 15: 0.0010328; 16: 0.0030166
| Rel.Std.Dev. |  | 01: 6.609%; 04: 1.52e-14%; 05: 2.236%; 06: 1.744%; 09: 8.155%; 10: 3.576%; 11: 2.875%; 12: 5.821%; 15: 2.312%; 16: 6.487%
| PDM3 |  | 01: -5.421%; 04: 12.15%; 05: -10.28%; 06: -3.178%; 09: 0.5607%; 10: 2.43%; 11: -1.308%; 12: 2.43%; 15: 0.1869%; 16: 4.299%
| Z-Score (Absolute) |  | 01: 0.92; 02: 7.23; 04: 1.96; 05: 1.72; 06: 0.55; 09: 0.06; 10: 0.37; 11: 0.25; 12: 0.37; 13: 1.96; 14: 7.23; 15: 0; 16: 0.67
| Analytical results for S in OREAS 554 (Certified Value 0.743 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 23 | 24 | 28 | 29 | [6 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [6 empty] | 3
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | 0.85 | 0.71 | 0.77 | 0.73 | 0.7735 | 0.76 | 0.792 | 0.78 | 0.72 | 0.64 | 0.71 | 0.68 | 0.71 | 0.5049 | 0.72 | 0.74 | 0.75 | 0.77 | 0.79 | [6 empty] | 1
| 1 | 2 | 0.84 | 0.72 | 0.76 | 0.72 | 0.7625 | 0.74 | 0.783 | 0.78 | 0.72 | 0.65 | 0.71 | 0.68 | 0.71 | 0.5044 | 0.73 | 0.74 | 0.77 | 0.8 | 0.8 | [6 empty] | 0
| 1 | 3 | 0.85 | 0.7 | 0.75 | 0.71 | 0.777 | 0.75 | 0.791 | 0.77 | 0.73 | 0.64 | 0.7 | 0.69 | 0.71 | 0.5238 | 0.81 | 0.73 | 0.77 | 0.77 | 0.8 | [6 empty] | 16
| 1 | 4 | 0.8 | 0.7 | 0.75 | 0.72 | 0.781 | 0.75 | 0.801 | 0.77 | 0.71 | 0.62 | 0.69 | 0.66 | 0.71 | 0.5118 | 0.73 | 0.73 | 0.77 | 0.8 | 0.79 | [6 empty] | 0.74281
| 1 | 5 | 0.84 | 0.69 | 0.77 | 0.73 | 0.766 | 0.76 | 0.788 | 0.77 | 0.7 | 0.62 | 0.7 | 0.65 | 0.72 | 0.5065 | 0.75 | 0.74 | 0.72 | 0.79 | 0.79 | [6 empty] | 181
| 1 | 6 | 0.85 | 0.72 | 0.76 | 0.74 | 0.7795 | 0.76 | 0.799 | 0.78 | 0.72 | 0.63 | 0.7 | 0.68 | 0.72 | 0.5063 | 0.74 | 0.75 | 0.71 | 0.78 | 0.8
| Mean |  | 0.83833 | 0.70667 | 0.76 | 0.725 | 0.77325 | 0.75333 | 0.79233 | 0.775 | 0.71667 | 0.63333 | 0.70167 | 0.67333 | 0.71333 | 0.50962 | 0.74667 | 0.73833 | 0.74833 | 0.785 | 0.795
| Median |  | 0.845 | 0.705 | 0.76 | 0.725 | 0.77525 | 0.755 | 0.7915 | 0.775 | 0.72 | 0.635 | 0.7 | 0.68 | 0.71 | 0.5064 | 0.735 | 0.74 | 0.76 | 0.785 | 0.795
| Std Dev. |  | 0.019408 | 0.012111 | 0.0089443 | 0.010488 | 0.0075017 | 0.008165 | 0.0067429 | 0.0054772 | 0.010328 | 0.012111 | 0.0075277 | 0.015055 | 0.005164 | 0.0074311 | 0.03266 | 0.0075277 | 0.027142 | 0.013784 | 0.0054772
| Rel.Std.Dev. |  | 2.315% | 1.714% | 1.177% | 1.447% | 0.9701% | 1.084% | 0.851% | 0.7067% | 1.441% | 1.912% | 1.073% | 2.236% | 0.7239% | 1.458% | 4.374% | 1.02% | 3.627% | 1.756% | 0.689%
| PDM3 |  | 12.86% | -4.866% | 2.314% | -2.398% | 4.098% | 1.417% | 6.667% | 4.334% | -3.52% | -14.74% | -5.539% | -9.353% | -3.968% | -31.39% | 0.5192% | -0.6027% | 0.7435% | 5.68% | 7.026%
| Z-Score (Absolute) |  | 1.85 | 0.81 | 0.27 | 0.44 | 0.54 | 0.13 | 0.92 | 0.57 | 0.61 | 2.29 | 0.91 | 1.48 | 0.67 | 4.8 | 0 | 0.17 | 0.03 | 0.78 | 0.98
| Analytical results for Sb in OREAS 554 (Certified Value 1.49 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 2.36 | 1.85 | 1.6 | 1.41 | 2.5 | 1.8 | 1.57 | 1.92 | < 5 | 1.19 | 1.09 | 1.46 | 1.35 | 1.6 | 0.62 | 1.2524 | 1.52 | 1.7597 | 1.58 | 1.4 | 1.57 | 1.58 | 0.37 |  |  | 1
| 1 | 2 | 2.33 | 1.88 | 1.5 | 1.52 | 2.5 | 1.6 | 1.6 | 1.83 | < 5 | 1.22 | 1.07 | 1.46 | 1.43 | 1.42 | 0.68 | 1.5985 | 1.43 | 1.7455 | 1.57 | 1.57 | 1.6 | 1.58 | 0.29 |  |  | 0
| 1 | 3 | 2.14 | 1.89 | 1.6 | 1.5 | 2.5 | 1.6 | 1.55 | 1.95 | < 5 | 1.23 | 1.08 | 1.42 | 1.46 | 1.51 | 0.79 | 1.389 | 1.48 | 1.7375 | 1.58 | 1.38 | 1.56 | 1.56 | 0.39 |  |  | 16
| 1 | 4 | 2.11 | 1.77 | 1.6 | 1.45 | 2.5 | 1.7 | 1.55 | 1.9 | < 5 | 1.18 | 1.14 | 1.42 | 1.42 | 1.26 | 0.73 | 1.3708 | 1.62 | 1.7583 | 1.53 | 1.53 | 1.6 | 1.54 | 0.34 |  |  | 1.4854
| 1 | 5 | 2.13 | 1.78 | 1.6 | 1.52 | 2.5 | 1.7 | 1.61 | 1.88 | < 5 | 1.16 | 1.16 | 1.39 | 1.45 | 1.21 | 0.71 | 1.6666 | 1.51 | 1.7236 | 1.54 | 1.42 | 1.6 | 1.58 | 0.32 |  |  | 182
| 1 | 6 | 2.15 | 1.82 | 1.6 | 1.45 | 2.5 | 1.7 | 1.5 | 1.95 | < 5 | 1.21 | 1.11 | 1.38 | 1.47 | 1.21 | 0.63 | 1.464 | 1.49 | 1.7522 | 1.57 | 1.5 | 1.54 | 1.56 | 0.32
| Mean |  | 2.2033 | 1.8317 | 1.5833 | 1.475 | 2.5 | 1.6833 | 1.5633 | 1.905 |  | 1.1983 | 1.1083 | 1.4217 | 1.43 | 1.3683 | 0.69333 | 1.4569 | 1.5083 | 1.7461 | 1.5617 | 1.4667 | 1.5783 | 1.5667 | 0.33833
| Median |  | 2.145 | 1.835 | 1.6 | 1.475 | 2.5 | 1.7 | 1.56 | 1.91 |  | 1.2 | 1.1 | 1.42 | 1.44 | 1.34 | 0.695 | 1.4265 | 1.5 | 1.7489 | 1.57 | 1.46 | 1.585 | 1.57 | 0.33
| Std Dev. |  | 0.11094 | 0.050365 | 0.040825 | 0.045056 | 0 | 0.075277 | 0.039833 | 0.045935 |  | 0.026394 | 0.035449 | 0.033714 | 0.043359 | 0.1663 | 0.064083 | 0.15358 | 0.063061 | 0.013777 | 0.02137 | 0.077374 | 0.025626 | 0.01633 | 0.03656
| Rel.Std.Dev. |  | 5.035% | 2.75% | 2.578% | 3.055% | 0% | 4.472% | 2.548% | 2.411% |  | 2.203% | 3.198% | 2.371% | 3.032% | 12.15% | 9.243% | 10.54% | 4.181% | 0.789% | 1.368% | 5.275% | 1.624% | 1.042% | 10.81%
| PDM3 |  | 48.33% | 23.31% | 6.591% | -0.7018% | 68.3% | 13.32% | 5.245% | 28.25% |  | -19.33% | -25.39% | -4.292% | -3.731% | -7.883% | -53.32% | -1.921% | 1.542% | 17.55% | 5.133% | -1.263% | 6.255% | 5.469% | -77.22%
| Z-Score (Absolute) |  | 4.57 | 2.18 | NA | 0.11 | NA | NA | 0.46 | 2.65 | 6.48 | 1.88 | 2.46 | 0.45 | 0.4 | 0.79 | 5.13 | 0.22 | 0.11 | 1.63 | 0.45 | 0.16 | 0.56 | 0.48 | 7.41
Results from laboratories 4, 6 & 9 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 554 (Certified Value 2.23 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | 2.3 | 2.1 | 3 | 2 | 2 | 2.9 | 2.3 | 2.3 | < 5 | 2.1 | 2.3 | 2.4 | 2.3 | 2.2 | 2.89 | 1.8571 | 2.2 | 2 | 2 | 2.4 | 2.6 | 2.3 |  |  |  | 1
| 1 | 2 | 2.3 | 2.2 | 2 | 2 | 2 | 2.4 | 2.1 | 2.3 | < 5 | 2.1 | 2.2 | 2.5 | 2.3 | 2.1 | 2.83 | 2.36 | 2.1 | 2 | 2 | 2.5 | 2.4 | 2.2 |  |  |  | 0
| 1 | 3 | 2.3 | 2 | 2 | 2 | 2 | 2.5 | 2.2 | 2.3 | < 5 | 2 | 2.3 | 2.4 | 2.3 | 2.1 | 2.96 | 1.9893 | 2.2 | 2 | 2 | 2.3 | 2.3 | 2.2 |  |  |  | 16
| 1 | 4 | 2.2 | 2 | 2 | 2 | 2 | 2.5 | 2.4 | 2.4 | < 5 | 1.9 | 2.5 | 2.4 | 2.3 | 2.1 | 2.88 | 2.0214 | 2.2 | 2 | 2 | 2.4 | 2.4 | 2.1 |  |  |  | 2.2326
| 1 | 5 | 2.2 | 2 | 2 | 2 | 2 | 2.4 | 2.3 | 2.2 | < 5 | 1.9 | 2.3 | 2.4 | 2.6 | 2.1 | 2.92 | 2.1751 | 2.2 | 2 | 2 | 2.3 | 2.3 | 2.2 |  |  |  | 183
| 1 | 6 | 2.3 | 2.1 | 2 | 2 | 2 | 2.4 | 2.2 | 2.3 | < 5 | 2 | 2.3 | 2.4 | 2.4 | 2.1 | 2.84 | 2.1656 | 2.1 | 2 | 1.9 | 2.3 | 2.5 | 2.2
| Mean |  | 2.2667 | 2.0667 | 2.1667 | 2 | 2 | 2.5167 | 2.25 | 2.3 |  | 2 | 2.3167 | 2.4167 | 2.3667 | 2.1167 | 2.8867 | 2.0947 | 2.1667 | 2 | 1.9833 | 2.3667 | 2.4167 | 2.2
| Median |  | 2.3 | 2.05 | 2 | 2 | 2 | 2.45 | 2.25 | 2.3 |  | 2 | 2.3 | 2.4 | 2.3 | 2.1 | 2.885 | 2.0935 | 2.2 | 2 | 2 | 2.35 | 2.4 | 2.2
| Std Dev. |  | 0.05164 | 0.08165 | 0.40825 | 0 | 0 | 0.19408 | 0.10488 | 0.063246 |  | 0.089443 | 0.098319 | 0.040825 | 0.12111 | 0.040825 | 0.048854 | 0.17598 | 0.05164 | 0 | 0.040825 | 0.08165 | 0.1169 | 0.063246
| Rel.Std.Dev. |  | 2.278% | 3.951% | 18.84% | 0% | 0% | 7.712% | 4.661% | 2.75% |  | 4.472% | 4.244% | 1.689% | 5.117% | 1.929% | 1.692% | 8.401% | 2.383% | 0% | 2.058% | 3.45% | 4.837% | 2.875%
| PDM3 |  | 1.526% | -7.432% | -2.953% | -10.42% | -10.42% | 12.72% | 0.7799% | 3.019% |  | -10.42% | 3.766% | 8.245% | 6.006% | -5.192% | 29.3% | -6.174% | -2.953% | -10.42% | -11.16% | 6.006% | 8.245% | -1.46%
| Z-Score (Absolute) |  | 0.08 | 1.1 | NA | NA | NA | 1.18 | 0.17 | 0.08 | 1.1 | 1.43 | 0.17 | 0.67 | 0.42 | 0.84 | 3.05 | 0.95 | 0.59 | NA | 1.52 | 0.42 | 0.67 | 0.42
Results from laboratories 4, 5, 6 & 33 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 554 (Certified Value 4.9 ppm) | [26 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 2
|  |  | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | 5 | 6 | 5 | 7 | 4 | 4.8 | 6.6 | < 5 | 4.6 | 5.4 | 5.3 | 5.3 | 5.2 | 3 | 4.891 | 5 | 4.5037 | 4 | 4.5 | 4.8 | < 10 | 2.8 |  |  |  | 1
| 1 | 2 | 4 | 6 | 4 | 6 | 3 | 5.2 | 6 | < 5 | 5.5 | 4.6 | 5.8 | 5.7 | 5.7 | 2.4 | 5.3641 | 5 | 4.1063 | 5 | 4.7 | 4.6 | < 10 | 3.2 |  |  |  | 0
| 1 | 3 | 5 | 6 | 5 | 7 | 4 | 4.6 | 6.1 | < 5 | 4.8 | 5.6 | 5.6 | 5.7 | 4.8 | 3.6 | 4.3247 | 5 | 4.5971 | 5 | 4.7 | 5.1 | < 10 | 3.4 |  |  |  | 16
| 1 | 4 | 5 | 6 | 4 | 7 | 4 | 4.8 | 6.4 | < 5 | 5 | 4.6 | 5.5 | 5.1 | 4.8 | 3 | 4.7937 | 5 | 4.3343 | 4 | 4.5 | 4.7 | < 10 | 3.9 |  |  |  | 4.8965
| 1 | 5 | 5 | 6 | 4 | 5 | 3 | 5.2 | 6.7 | < 5 | 5.1 | 4.3 | 5.4 | 5 | 4.7 | 2.9 | 5.1984 | 5 | 4.7529 | 5 | 4.7 | 4.6 | < 10 | 4.2 |  |  |  | 184
| 1 | 6 | 4 | 6 | 4 | 7 | 4 | 4.9 | 6.1 | < 5 | 4.7 | 5.2 | 5.9 | 5.6 | 5.3 | 2.9 | 4.8415 | 6 | 5.1972 | 5 | 4.8 | 4.4 | < 10 | 3.6
| Mean |  | 4.6667 | 6 | 4.3333 | 6.5 | 3.6667 | 4.9167 | 6.3167 |  | 4.95 | 4.95 | 5.5833 | 5.4 | 5.0833 | 2.9667 | 4.9022 | 5.1667 | 4.5819 | 4.6667 | 4.65 | 4.7 |  | 3.5167
| Median |  | 5 | 6 | 4 | 7 | 4 | 4.85 | 6.25 |  | 4.9 | 4.9 | 5.55 | 5.45 | 5 | 2.95 | 4.8663 | 5 | 4.5504 | 5 | 4.7 | 4.65 |  | 3.5
| Std Dev. |  | 0.5164 | 0 | 0.5164 | 0.83666 | 0.5164 | 0.24014 | 0.29269 |  | 0.32711 | 0.52058 | 0.23166 | 0.30984 | 0.38687 | 0.38297 | 0.3606 | 0.40825 | 0.37462 | 0.5164 | 0.12247 | 0.23664 |  | 0.49967
| Rel.Std.Dev. |  | 11.07% | 0% | 11.92% | 12.87% | 14.08% | 4.884% | 4.634% |  | 6.608% | 10.52% | 4.149% | 5.738% | 7.611% | 12.91% | 7.356% | 7.902% | 8.176% | 11.07% | 2.634% | 5.035% |  | 14.21%
| PDM3 |  | -4.694% | 22.54% | -11.5% | 32.75% | -25.12% | 0.4121% | 29% |  | 1.093% | 1.093% | 14.03% | 10.28% | 3.816% | -39.41% | 0.1174% | 5.518% | -6.424% | -4.694% | -5.034% | -4.013% |  | -28.18%
| Z-Score (Absolute) |  | 0.56 | 2.51 | 1.32 | 3.65 | 2.86 | 0.02 | 3.23 | 5.54 | 0.09 | 0.09 | 1.55 | 1.13 | 0.4 | 4.46 | 0.02 | 0.59 | 0.75 | 0.56 | 0.6 | 0.48 | 0.21 | 3.2
| Analytical results for Si in OREAS 554 (Indicative Value 0.061 wt.%) | [26 empty] | Indicative
Si | Round | Replicate | 02: Lab |  |  |  | 1
| No | No | 02: 06 |  |  |  | wt.%
|  |  | 02: AR*MS |  |  |  | 3
|  |  | 02: 25g |  |  |  | 3
| 1 | 1 | 02: 0.0605 |  |  |  | 1
| 1 | 2 | 02: 0.06 |  |  |  | 0
| 1 | 3 | 02: 0.0605 |  |  |  | 16
| 1 | 4 | 02: 0.061 |  |  |  | 0.061083
| 1 | 5 | 02: 0.06 |  |  |  | 29
| 1 | 6 | 02: 0.0645
| Mean |  | 02: 0.061083
| Median |  | 02: 0.0605
| Std Dev. |  | 02: 0.0017151
| Rel.Std.Dev. |  | 02: 2.808%
| PDM3 |  | 02: 4.441e-14%
Z-Score (Absolute) | bar: 02
| Analytical results for Sm in OREAS 554 (Indicative Value 2.51 ppm) | [26 empty] | Indicative
Sm | Round | Replicate | 02: Lab; 04: Lab; 05: Lab; 06: Lab; 09: Lab |  |  |  | 1
| No | No | 02: 02; 04: 04; 05: 06; 06: 10; 09: 30 |  |  |  | ppm
|  |  | 02: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 09: AR*MS |  |  |  | 2
|  |  | 02: 25g; 04: 01-05g; 05: 25g; 06: 0.5g; 09: 0.5g |  |  |  | 2
| 1 | 1 | 02: 2.679; 04: 2.25; 05: 2.9; 06: 3.4; 09: 1.66 |  |  |  | 1
| 1 | 2 | 02: 2.85; 04: 2.18; 05: 2.7; 06: 3.1; 09: 1.64 |  |  |  | 0
| 1 | 3 | 02: 2.742; 04: 2.16; 05: 2.7; 06: 3.1; 09: 1.7 |  |  |  | 16
| 1 | 4 | 02: 2.635; 04: 2.11; 05: 2.8; 06: 3.3; 09: 1.7 |  |  |  | 2.5115
| 1 | 5 | 02: 2.678; 04: 2.12; 05: 2.6; 06: 3.3; 09: 1.66 |  |  |  | 30
| 1 | 6 | 02: 2.75; 04: 2.18; 05: 2.9; 06: 3.1; 09: 1.75
| Mean |  | 02: 2.7223; 04: 2.1667; 05: 2.7667; 06: 3.2167; 09: 1.685
| Median |  | 02: 2.7105; 04: 2.17; 05: 2.75; 06: 3.2; 09: 1.68
| Std Dev. |  | 02: 0.076054; 04: 0.050465; 05: 0.12111; 06: 0.13292; 09: 0.039875
| Rel.Std.Dev. |  | 02: 2.794%; 04: 2.329%; 05: 4.377%; 06: 4.132%; 09: 2.366%
| PDM3 |  | 02: 8.396%; 04: -13.73%; 05: 10.16%; 06: 28.08%; 09: -32.91%
| Z-Score (Absolute) |  | 02: 0; 04: 0.76; 05: 0.06; 06: 0.67; 09: 1.42
| Analytical results for Sn in OREAS 554 (Certified Value 0.32 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.25g | 0.5g |  |  |  |  | 2
| 1 | 1 | 0.4 | 0.34 | < 1 | < 0.5 | < 0.5 | 0.39 | 0.5 | < 10 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.2 | 0.63015 | 0.4 | 0.30586 | 1.1 | 0.4 | < 0.3 | 0.22 |  |  |  |  | 1
| 1 | 2 | 0.4 | 0.33 | < 1 | < 0.5 | < 0.5 | 0.38 | 0.5 | < 10 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.2 | 0.50578 | 0.4 | 0.29346 | 1.1 | 0.4 | < 0.3 | 0.24 |  |  |  |  | 0
| 1 | 3 | 0.3 | 0.33 | < 1 | < 0.5 | < 0.5 | 0.38 | 0.6 | < 10 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.83544 | 0.4 | 0.28927 | 1 | 0.4 | < 0.3 | 0.22 |  |  |  |  | 16
| 1 | 4 | 0.3 | 0.32 | < 1 | < 0.5 | < 0.5 | 0.39 | 0.5 | < 10 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.71612 | 1.7 | 0.29536 | 1 | 0.4 | < 0.3 | 0.25 |  |  |  |  | 0.31905
| 1 | 5 | 0.3 | 0.33 | < 1 | < 0.5 | < 0.5 | 0.4 | 0.5 | < 10 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.88151 | 0.4 | 0.30501 | 1 | 0.4 | < 0.3 | 0.22 |  |  |  |  | 185
| 1 | 6 | 0.3 | 0.34 | < 1 | < 0.5 | < 0.5 | 0.36 | 0.5 | < 10 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.95503 | 0.4 | 0.31678 | 1 | 0.4 | < 0.3 | 0.24
| Mean |  | 0.33333 | 0.33167 |  |  |  | 0.38333 | 0.51667 |  | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.26667 | 0.75401 | 0.61667 | 0.30096 | 1.0333 | 0.4 |  | 0.23167
| Median |  | 0.3 | 0.33 |  |  |  | 0.385 | 0.5 |  | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.3 | 0.77578 | 0.4 | 0.30018 | 1 | 0.4 |  | 0.23
| Std Dev. |  | 0.05164 | 0.0075277 |  |  |  | 0.013663 | 0.040825 |  | 0 | 0 | 0 | 0 | 0 | 0.05164 | 0.16834 | 0.53072 | 0.010148 | 0.05164 | 6.0809e-17 |  | 0.013292
| Rel.Std.Dev. |  | 15.49% | 2.27% |  |  |  | 3.564% | 7.902% |  | 0% | 0% | 0% | 0% | 0% | 19.36% | 22.33% | 86.06% | 3.372% | 4.997% | 1.52e-14% |  | 5.737%
| PDM3 |  | 4.478% | 3.955% |  |  |  | 20.15% | 61.94% |  | -5.97% | -5.97% | -5.97% | -5.97% | -5.97% | -16.42% | 136.3% | 93.28% | -5.67% | 223.9% | 25.37% |  | -27.39%
| Z-Score (Absolute) |  | 0.43 | 0.41 | 2.63 | 0.67 | 0.67 | 1.09 | 2.85 | 62.18 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | 0.45 | 6 | 4.18 | 0 | 9.69 | 1.31 | 2 | 0.92
| Analytical results for Sr in OREAS 554 (Certified Value 40.2 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g |  |  |  |  | 2
| 1 | 1 | 39.7 | 41.01 | 43 | 36.5 | 39 | 43 | 41.6 | 43.6 | 39 | 42.5 | 41.3 | 42.8 | 41.4 | 40.5 | 41 | 35.3884 | 37.8 | 41.5 | 36 | 41.6 | 40.9 |  |  |  |  | 1
| 1 | 2 | 40.5 | 42.23 | 44 | 35.9 | 38 | 43 | 41.6 | 43.3 | 38 | 43 | 40.2 | 43.6 | 40.7 | 39.1 | 41 | 36.025 | 37.7 | 41.5 | 37.5 | 42 | 41.1 |  |  |  |  | 0
| 1 | 3 | 39.7 | 41.03 | 41 | 39.4 | 38 | 43 | 42.6 | 45.9 | 39 | 43.5 | 41.2 | 42.5 | 40.8 | 39.5 | 48 | 32.3393 | 38.4 | 41.8 | 37.2 | 40.2 | 40.6 |  |  |  |  | 16
| 1 | 4 | 37.7 | 39.94 | 41 | 38.4 | 38 | 42 | 43.2 | 43.5 | 39 | 42.1 | 40.4 | 42.3 | 40.2 | 39.4 | 44 | 33.6929 | 38.3 | 40.8 | 36.4 | 41.8 | 40.6 |  |  |  |  | 40.2326
| 1 | 5 | 38.1 | 39.6 | 40 | 38.9 | 37 | 43 | 42.8 | 43 | 39 | 41.3 | 38.9 | 41.7 | 36.5 | 39.4 | 42 | 33.7176 | 39 | 40.4 | 35.5 | 40.4 | 40.8 |  |  |  |  | 186
| 1 | 6 | 38.3 | 41.38 | 41 | 37.3 | 39 | 43 | 41.1 | 44.1 | 39 | 43.1 | 39.4 | 42.4 | 40.3 | 41.4 | 42 | 35.0624 | 38.6 | 40.8 | 34.3 | 42.1 | 39.4
| Mean |  | 39 | 40.865 | 41.6667 | 37.7333 | 38.1667 | 42.8333 | 42.15 | 43.9 | 38.8333 | 42.5833 | 40.2333 | 42.55 | 39.9833 | 39.8833 | 43 | 34.3709 | 38.3 | 41.1333 | 36.15 | 41.35 | 40.5667
| Median |  | 39 | 41.02 | 41 | 37.85 | 38 | 43 | 42.1 | 43.55 | 39 | 42.75 | 40.3 | 42.45 | 40.5 | 39.45 | 42 | 34.39 | 38.35 | 41.15 | 36.2 | 41.7 | 40.7
| Std Dev. |  | 1.1153 | 0.96255 | 1.5055 | 1.3895 | 0.75277 | 0.40825 | 0.82885 | 1.045 | 0.40825 | 0.79603 | 0.95638 | 0.62849 | 1.7589 | 0.88412 | 2.6833 | 1.3611 | 0.4899 | 0.54283 | 1.1709 | 0.83367 | 0.60222
| Rel.Std.Dev. |  | 2.86% | 2.355% | 3.613% | 3.682% | 1.972% | 0.9531% | 1.966% | 2.38% | 1.051% | 1.869% | 2.377% | 1.477% | 4.399% | 2.217% | 6.24% | 3.96% | 1.279% | 1.32% | 3.239% | 2.016% | 1.485%
| PDM3 |  | -3.064% | 1.572% | 3.564% | -6.212% | -5.135% | 6.464% | 4.766% | 9.115% | -3.478% | 5.843% | 0.001705% | 5.76% | -0.6197% | -0.8682% | 6.878% | -14.57% | -4.804% | 2.239% | -10.15% | 2.777% | 0.8302%
| Z-Score (Absolute) |  | 0.61 | 0.12 | 0.43 | 1.1 | 0.93 | 0.88 | 0.62 | 1.3 | 0.67 | 0.78 | 0.13 | 0.77 | 0.23 | 0.27 | 0.95 | 2.41 | 0.88 | 0.22 | 1.72 | 0.3 | 0
| Analytical results for Ta in OREAS 554 (Certified Value < 0.01 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 3
|  |  | 25g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [7 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 0.5 | < 0.05 | 0.08 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | NA | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | 0.01 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 0.5 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | NA | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | [7 empty] | 0
| 1 | 3 | < 0.01 | < 0.05 | < 0.5 | < 0.05 | 0.08 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | NA | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | 0.01 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 0.5 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | NA | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | 0.01 | [7 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 0.5 | < 0.05 | < 0.05 | < 5 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | NA | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | 0.01 | [7 empty] | 187
| 1 | 6 | < 0.01 | < 0.05 | < 0.5 | < 0.05 | < 0.05 | < 5 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | NA | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.05 | < 0.01
| Mean |  | 11: 0.08; 14: 0.01; 30: 0.01
| Median |  | 11: 0.08; 14: 0.01; 30: 0.01
| Std Dev. |  | 11: 0; 30: 0
Rel.Std.Dev. | bar: 11; 30
PDM3
| Z-Score (Absolute) |  | 0.67 | 3.37 | 48.89 | 3.37 | 7.08 | 504.05 | 0.67 | 0.51 | 0.67 | 0.67 | 0.67 | NA | 0.67 | 3.37 | 3.37 | 0.67 | 3.37 | 0
| Analytical results for Tb in OREAS 554 (Certified Value 0.29 ppm) | [26 empty] | Certified
Tb | Round | Replicate | 02: Lab; 05: Lab; 06: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | [7 empty] | 1
| No | No | 02: 02; 05: 04; 06: 06; 10: 10; 11: 11; 12: 21; 13: 29; 14: 30 | [7 empty] | ppm
|  |  | 02: AR*MS; 05: AR*MS; 06: AR*MS; 10: AR*OES/MS; 11: AR*OES/MS; 12: AR*OES/MS; 13: AR*OES/MS; 14: AR*MS | [7 empty] | 2
|  |  | 02: 25g; 05: 01-05g; 06: 25g; 10: 0.5g; 11: 0.25g; 12: 0.25g; 13: 0.25g; 14: 0.5g | [7 empty] | 3
| 1 | 1 | 02: 0.301; 05: 0.26; 06: 0.3; 10: 0.3; 11: 0.42; 12: 0.28; 13: 0.31; 14: 0.22 | [7 empty] | 1
| 1 | 2 | 02: 0.309; 05: 0.255; 06: 0.3; 10: 0.3; 11: 0.4; 12: 0.28; 13: 0.32; 14: 0.24 | [7 empty] | 0
| 1 | 3 | 02: 0.294; 05: 0.245; 06: 0.3; 10: 0.3; 11: 0.46; 12: 0.29; 13: 0.32; 14: 0.25 | [7 empty] | 16
| 1 | 4 | 02: 0.28; 05: 0.245; 06: 0.3; 10: 0.3; 11: 0.41; 12: 0.3; 13: 0.33; 14: 0.22 | [7 empty] | 0.29164
| 1 | 5 | 02: 0.292; 05: 0.24; 06: 0.3; 10: 0.3; 11: 0.41; 12: 0.29; 13: 0.31; 14: 0.23 | [7 empty] | 188
| 1 | 6 | 02: 0.298; 05: 0.25; 06: 0.3; 10: 0.3; 11: 0.42; 12: 0.28; 13: 0.32; 14: 0.24
| Mean |  | 02: 0.29567; 05: 0.24917; 06: 0.3; 10: 0.3; 11: 0.42; 12: 0.28667; 13: 0.31833; 14: 0.23333
| Median |  | 02: 0.296; 05: 0.2475; 06: 0.3; 10: 0.3; 11: 0.415; 12: 0.285; 13: 0.32; 14: 0.235
| Std Dev. |  | 02: 0.0097297; 05: 0.0073598; 06: 0; 10: 0; 11: 0.020976; 12: 0.008165; 13: 0.0075277; 14: 0.012111
| Rel.Std.Dev. |  | 02: 3.291%; 05: 2.954%; 06: 0%; 10: 0%; 11: 4.994%; 12: 2.848%; 13: 2.365%; 14: 5.19%
| PDM3 |  | 02: 1.381%; 05: -14.56%; 06: 2.867%; 10: 2.867%; 11: 44.01%; 12: -1.705%; 13: 9.153%; 14: -19.99%
| Z-Score (Absolute) |  | 02: 0.09; 05: 2.07; 06: 0.09; 10: 0.09; 11: 5.2; 12: 0.48; 13: 0.87; 14: 2.75
| Analytical results for Te in OREAS 554 (Certified Value 0.048 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  |  | 3
| 1 | 1 | < 0.1 | < 0.2 | 0.06 | 0.1 | 0.05 | 0.39 | < 5 | 0.04 | 0.05 | 0.05 | 0.04 | 0.06 | 0.04 | < 0.01 | < 0.05 | 0.04288 | 0.06 | 0.04 | < 0.01 | < 0.05 | 0.05 |  |  |  |  | 1
| 1 | 2 | < 0.1 | < 0.2 | 0.07 | 0.1 | 0.06 | 0.26 | < 5 | 0.05 | 0.04 | 0.05 | 0.05 | 0.07 | 0.03 | < 0.01 | < 0.05 | 0.04767 | 0.06 | 0.05 | 0.02 | < 0.05 | 0.03 |  |  |  |  | 0
| 1 | 3 | < 0.1 | < 0.2 | 0.07 | 0.1 | 0.06 | 0.29 | < 5 | 0.05 | 0.05 | 0.05 | 0.06 | 0.03 | 0.04 | < 0.01 | 0.05 | 0.05013 | 0.06 | < 0.02 | 0.07 | < 0.05 | 0.02 |  |  |  |  | 16
| 1 | 4 | < 0.1 | < 0.2 | 0.05 | 0.1 | 0.03 | 0.26 | < 5 | 0.04 | 0.04 | 0.05 | 0.04 | 0.05 | 0.03 | < 0.01 | < 0.05 | 0.04843 | 0.06 | 0.05 | 0.09 | < 0.05 | 0.05 |  |  |  |  | 0.048241
| 1 | 5 | < 0.1 | < 0.2 | 0.06 | 0.1 | 0.04 | 0.18 | < 5 | 0.04 | 0.05 | 0.05 | 0.05 | 0.04 | 0.03 | < 0.01 | < 0.05 | 0.038 | 0.06 | 0.03 | 0.02 | < 0.05 | 0.04 |  |  |  |  | 189
| 1 | 6 | < 0.1 | < 0.2 | 0.06 | 0.1 | 0.05 | 0.16 | < 5 | 0.04 | 0.04 | 0.06 | 0.05 | 0.03 | 0.03 | < 0.01 | 0.06 | 0.05082 | 0.06 | 0.06 | 0.05 | < 0.05 | 0.01
| Mean |  |  |  | 0.061667 | 0.1 | 0.048333 | 0.25667 |  | 0.043333 | 0.045 | 0.051667 | 0.048333 | 0.046667 | 0.033333 |  | 0.055 | 0.046322 | 0.06 | 0.046 | 0.05 |  | 0.033333
| Median |  |  |  | 0.06 | 0.1 | 0.05 | 0.26 |  | 0.04 | 0.045 | 0.05 | 0.05 | 0.045 | 0.03 |  | 0.055 | 0.04805 | 0.06 | 0.05 | 0.05 |  | 0.035
| Std Dev. |  |  |  | 0.0075277 | 1.5202e-17 | 0.01169 | 0.082624 |  | 0.005164 | 0.0054772 | 0.0040825 | 0.0075277 | 0.01633 | 0.005164 |  | 0.0070711 | 0.0049416 | 0 | 0.011402 | 0.030822 |  | 0.01633
| Rel.Std.Dev. |  |  |  | 12.21% | 1.52e-14% | 24.19% | 32.19% |  | 11.92% | 12.17% | 7.902% | 15.57% | 34.99% | 15.49% |  | 12.86% | 10.67% | 0% | 24.79% | 61.64% |  | 48.99%
| PDM3 |  |  |  | 27.83% | 107.3% | 0.1906% | 432% |  | -10.17% | -6.719% | 7.1% | 0.1906% | -3.264% | -30.9% |  | 14.01% | -3.979% | 24.37% | -4.646% | 3.645% |  | -30.9%
| Z-Score (Absolute) |  | 0.26 | 4.01 | 1.14 | NA | 0.14 | 15.76 | 183.94 | 0.24 | 0.11 | 0.39 | 0.14 | 0.01 | 0.99 | 3.11 | 0.86 | 0.01 | 1.01 | 0.49 | 0.3 | 1.61 | 0.99
Results from laboratory 6 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 554 (Certified Value 5.51 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  |  | 3
| 1 | 1 | 5.1 | 5.64 | 4.95 | 5.39 | 5.2 | 5.5 | 6.3 | 5.4 | 5.4 | 5.5 | 6.1 | 5.5 | 5.4 | 5.5 | 4.8105 | 4.3 | 5.64 | 5.7 | 6.1 | 5.4 | 4.31 |  |  |  |  | 1
| 1 | 2 | 5.2 | 5.71 | 4.95 | 5.33 | 5 | 5.3 | 6.3 | 5.3 | 5.4 | 5.6 | 6 | 5.7 | 5.4 | 5.5 | 5.2408 | 4.3 | 5.54 | 5.7 | 5.9 | 5.6 | 4.29 |  |  |  |  | 0
| 1 | 3 | 5.8 | 5.49 | 5.05 | 5.53 | 5 | 5.6 | 6.2 | 5.5 | 5.5 | 5.5 | 5.8 | 5.8 | 5.4 | 6.2 | 4.5584 | 4.5 | 5.45 | 5.8 | 5.9 | 5.4 | 4.46 |  |  |  |  | 16
| 1 | 4 | 5.7 | 5.45 | 5.1 | 5.61 | 5.1 | 5.7 | 6.4 | 5.2 | 5.3 | 5.6 | 5.8 | 5.7 | 5.3 | 5.7 | 4.6759 | 4.5 | 5.23 | 5.4 | 6.1 | 5.4 | 4.41 |  |  |  |  | 5.5111
| 1 | 5 | 5.7 | 5.46 | 5.2 | 5.7 | 5 | 5.8 | 6.4 | 5.3 | 5.4 | 5.4 | 5.8 | 5.5 | 5.5 | 5.5 | 4.8948 | 4.4 | 5.4 | 5.9 | 6 | 5.7 | 4.32 |  |  |  |  | 190
| 1 | 6 | 5.7 | 5.71 | 5.05 | 5.48 | 5.3 | 5.9 | 6.4 | 5.4 | 5.4 | 5.4 | 5.8 | 5.8 | 5.3 | 5.5 | 5.1693 | 4.5 | 5.39 | 5.7 | 5.8 | 5.5 | 4.31
| Mean |  | 5.5333 | 5.5767 | 5.05 | 5.5067 | 5.1 | 5.6333 | 6.3333 | 5.35 | 5.4 | 5.5 | 5.8833 | 5.6667 | 5.3833 | 5.65 | 4.8916 | 4.4167 | 5.4417 | 5.7 | 5.9667 | 5.5 | 4.35
| Median |  | 5.7 | 5.565 | 5.05 | 5.505 | 5.05 | 5.65 | 6.35 | 5.35 | 5.4 | 5.5 | 5.8 | 5.7 | 5.4 | 5.5 | 4.8527 | 4.45 | 5.425 | 5.7 | 5.95 | 5.45 | 4.315
| Std Dev. |  | 0.30111 | 0.12388 | 0.094868 | 0.13721 | 0.12649 | 0.21602 | 0.08165 | 0.10488 | 0.063246 | 0.089443 | 0.13292 | 0.13663 | 0.075277 | 0.28107 | 0.26951 | 0.098319 | 0.1402 | 0.16733 | 0.12111 | 0.12649 | 0.068411
| Rel.Std.Dev. |  | 5.442% | 2.221% | 1.879% | 2.492% | 2.48% | 3.835% | 1.289% | 1.96% | 1.171% | 1.626% | 2.259% | 2.411% | 1.398% | 4.975% | 5.51% | 2.226% | 2.576% | 2.936% | 2.03% | 2.3% | 1.573%
| PDM3 |  | 0.4038% | 1.19% | -8.366% | -0.08005% | -7.459% | 2.218% | 14.92% | -2.923% | -2.016% | -0.201% | 6.755% | 2.823% | -2.318% | 2.521% | -11.24% | -19.86% | -1.259% | 3.428% | 8.267% | -0.201% | -21.07%
| Z-Score (Absolute) |  | 0.15 | 0.34 | 2.02 | 0.03 | 1.8 | 0.6 | 3.75 | 0.67 | 0.45 | 0 | 1.72 | 0.75 | 0.52 | 0.67 | 2.73 | 4.87 | 0.26 | 0.9 | 2.1 | 0 | 5.17
| Analytical results for Ti in OREAS 554 (Indicative Value 0.003 wt.%) | [26 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 05 | 06 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 21 | 23 | 24 | 28 | 29 | [6 empty] | wt.%
|  |  | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | [6 empty] | 3
|  |  | 0.25g | 25g | 01g | 25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | [6 empty] | 3
| 1 | 1 | < 0.01 | 0.0022 | 0.0036 | 0.003 | 0.004 | < 0.01 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | NA | < 0.005 | < 0.01 | 0.0043 | 0.002 | < 0.005 | < 0.01 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.0023 | 0.0038 | 0.003 | 0.004 | < 0.01 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | NA | < 0.005 | < 0.01 | 0.0043 | 0.002 | < 0.005 | < 0.01 | [6 empty] | 0
| 1 | 3 | < 0.01 | 0.0022 | 0.0046 | 0.003 | 0.004 | < 0.01 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | NA | < 0.005 | < 0.01 | 0.0042 | 0.002 | < 0.005 | < 0.01 | [6 empty] | 16
| 1 | 4 | < 0.01 | 0.0021 | 0.0044 | 0.003 | 0.004 | < 0.01 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | NA | < 0.005 | < 0.01 | 0.0043 | 0.002 | < 0.005 | < 0.01 | [6 empty] | 0.0032556
| 1 | 5 | < 0.01 | 0.0021 | 0.0045 | 0.0025 | 0.004 | < 0.01 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | NA | < 0.005 | < 0.01 | 0.0042 | 0.002 | < 0.005 | < 0.01 | [6 empty] | 27
| 1 | 6 | < 0.01 | 0.0022 | 0.0037 | 0.003 | 0.004 | < 0.01 | < 0.01 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | < 0.005 | NA | < 0.005 | < 0.01 | 0.0042 | 0.002 | < 0.005 | < 0.01
| Mean |  | 02: 0.0021833; 04: 0.0041; 05: 0.0029167; 06: 0.004; 21: 0.00425; 23: 0.002
| Median |  | 02: 0.0022; 04: 0.0041; 05: 0.003; 06: 0.004; 21: 0.00425; 23: 0.002
| Std Dev. |  | 02: 7.5277e-05; 04: 0.00044721; 05: 0.00020412; 06: 0; 21: 5.4772e-05; 23: 0
| Rel.Std.Dev. |  | 02: 3.448%; 04: 10.91%; 05: 6.999%; 06: 0%; 21: 1.289%; 23: 0%
| PDM3 |  | 02: -32.94%; 04: 25.94%; 05: -10.41%; 06: 22.87%; 21: 30.55%; 23: -38.57%
| Z-Score (Absolute) |  | 2.51 | 0.57 | 1.52 | 0.23 | 1.41 | 2.51 | 2.51 | 0.23 | 0.23 | 0.23 | 0.23 | 0.23 | NA | 0.23 | 2.51 | 1.69 | 0.77 | 0.23 | 2.51
| Analytical results for Tl in OREAS 554 (Certified Value 0.11 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 | 23 | 24 | 28 | 29 |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g |  |  |  | 2
| 1 | 1 | 0.11 | 0.11 | 0.1 | 0.12 | 0.1 | < 0.5 | 0.11 | 0.13 | 0.1 | 0.11 | 0.09 | 0.11 | 0.1 | 0.077 | 0.11237 | < 5 | 0.07 | 0.11698 | 0.12 | 0.1 | 0.12 | 0.11 |  |  |  | 1
| 1 | 2 | 0.1 | 0.11 | 0.1 | 0.14 | 0.1 | < 0.5 | 0.11 | 0.12 | 0.11 | 0.11 | 0.09 | 0.12 | 0.09 | 0.075 | 0.13215 | < 5 | 0.07 | 0.10298 | 0.12 | 0.12 | 0.12 | 0.1 |  |  |  | 0
| 1 | 3 | 0.1 | 0.1 | < 0.1 | 0.13 | 0.1 | < 0.5 | 0.11 | 0.17 | 0.1 | 0.11 | 0.09 | 0.13 | 0.1 | 0.086 | 0.10843 | < 5 | 0.07 | 0.10967 | 0.13 | 0.11 | 0.12 | 0.11 |  |  |  | 16
| 1 | 4 | 0.1 | 0.11 | < 0.1 | 0.12 | 0.1 | < 0.5 | 0.11 | 0.12 | 0.1 | 0.11 | 0.09 | 0.11 | 0.1 | 0.079 | 0.10957 | < 5 | 0.08 | 0.0934 | 0.12 | 0.11 | 0.12 | 0.1 |  |  |  | 0.10886
| 1 | 5 | 0.1 | 0.11 | < 0.1 | 0.12 | 0.1 | < 0.5 | 0.12 | 0.12 | 0.1 | 0.12 | 0.09 | 0.12 | 0.1 | 0.075 | 0.1286 | < 5 | 0.08 | 0.11382 | 0.12 | 0.11 | 0.12 | 0.11 |  |  |  | 191
| 1 | 6 | 0.1 | 0.11 | < 0.1 | 0.12 | 0.1 | < 0.5 | 0.11 | 0.12 | 0.11 | 0.12 | 0.09 | 0.12 | 0.09 | 0.075 | 0.12692 | < 5 | 0.08 | 0.11013 | 0.12 | 0.11 | 0.11 | 0.11
| Mean |  | 0.10167 | 0.10833 | 0.1 | 0.125 | 0.1 |  | 0.11167 | 0.13 | 0.10333 | 0.11333 | 0.09 | 0.11833 | 0.096667 | 0.077833 | 0.11967 |  | 0.075 | 0.10783 | 0.12167 | 0.11 | 0.11833 | 0.10667
| Median |  | 0.1 | 0.11 | 0.1 | 0.12 | 0.1 |  | 0.11 | 0.12 | 0.1 | 0.11 | 0.09 | 0.12 | 0.1 | 0.076 | 0.11964 |  | 0.075 | 0.1099 | 0.12 | 0.11 | 0.12 | 0.11
| Std Dev. |  | 0.0040825 | 0.0040825 | 0 | 0.0083666 | 1.5202e-17 |  | 0.0040825 | 0.02 | 0.005164 | 0.005164 | 1.5202e-17 | 0.0075277 | 0.005164 | 0.0043089 | 0.010674 |  | 0.0054772 | 0.0084856 | 0.0040825 | 0.0063246 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 4.016% | 3.768% | 0% | 6.693% | 1.52e-14% |  | 3.656% | 15.38% | 4.997% | 4.556% | 1.689e-14% | 6.361% | 5.342% | 5.536% | 8.919% |  | 7.303% | 7.869% | 3.355% | 5.75% | 3.45% | 4.841%
| PDM3 |  | -6.612% | -0.4882% | -8.143% | 14.82% | -8.143% |  | 2.574% | 19.41% | -5.081% | 4.105% | -17.33% | 8.697% | -11.2% | -28.5% | 9.927% |  | -31.11% | -0.9506% | 11.76% | 1.043% | 8.697% | -2.019%
| Z-Score (Absolute) |  | 0.69 | 0.19 | NA | 1.06 | NA | 10.42 | 0.06 | 1.44 | 0.56 | 0.19 | 1.56 | 0.56 | 1.06 | 2.47 | 0.66 | 178.95 | 2.68 | 0.22 | 0.81 | 0.06 | 0.56 | 0.31
Results from laboratories 4 & 6 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 554 (Indicative Value 0.086 ppm) | [26 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab |  |  |  | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS |  |  |  | 3
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g |  |  |  | 3
| 1 | 1 | 01: 0.079; 02: 0.08; 04: 0.1; 05: < 0.1 |  |  |  | 1
| 1 | 2 | 01: 0.082; 02: 0.075; 04: 0.1; 05: < 0.1 |  |  |  | 0
| 1 | 3 | 01: 0.084; 02: 0.075; 04: 0.1; 05: < 0.1 |  |  |  | 16
| 1 | 4 | 01: 0.08; 02: 0.075; 04: 0.1; 05: < 0.1 |  |  |  | 0.085667
| 1 | 5 | 01: 0.078; 02: 0.075; 04: 0.05; 05: < 0.1 |  |  |  | 28
| 1 | 6 | 01: 0.084; 02: 0.075; 04: 0.1; 05: < 0.1
| Mean |  | 01: 0.081167; 02: 0.075833; 04: 0.091667
| Median |  | 01: 0.081; 02: 0.075; 04: 0.1
| Std Dev. |  | 01: 0.0025626; 02: 0.0020412; 04: 0.020412
| Rel.Std.Dev. |  | 01: 3.157%; 02: 2.692%; 04: 22.27%
| PDM3 |  | 01: -5.253%; 02: -11.48%; 04: 7.004%
| Z-Score (Absolute) |  | 01: 0.23; 02: 0.23; 04: 1.12; 05: 2.43
| Analytical results for U in OREAS 554 (Certified Value 11.8 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  |  | 2
| 1 | 1 | 11.26 | 12.07 | 11.6 | 11.76 | 11.5 | 12 | 10.6 | 14.1 | 12.5 | 11.7 | 12.95 | 11.9 | 12.5 | 14.8 | 11.3799 | 12.44 | 11.2 | 11.06 | 13 | 11.97 | 12.1 |  |  |  |  | 1
| 1 | 2 | 11.02 | 12 | 11 | 11.67 | 11 | 11.7 | 10.6 | 13.8 | 12.1 | 11.7 | 12.55 | 12.1 | 14 | 15.1 | 11.7851 | 11.83 | 11 | 11.19 | 13.17 | 12.55 | 11.4 |  |  |  |  | 0
| 1 | 3 | 10.02 | 11.82 | 11 | 11.71 | 11 | 12.4 | 10.4 | 13.7 | 12.8 | 11.95 | 12.5 | 12.05 | 13.15 | 16.9 | 10.2793 | 11.9 | 11 | 11.45 | 12.57 | 12.73 | 11.7 |  |  |  |  | 16
| 1 | 4 | 10.2 | 11.72 | 11.3 | 11.85 | 11.3 | 12.1 | 10.6 | 13.6 | 12.2 | 12.05 | 12.3 | 12.15 | 11.55 | 16.2 | 11.056 | 11.86 | 11.2 | 10.6 | 12.89 | 12 | 12 |  |  |  |  | 11.8427
| 1 | 5 | 10.2 | 11.99 | 11.4 | 11.86 | 11.2 | 12.6 | 10.7 | 13.8 | 12.05 | 11.65 | 12.8 | 12.1 | 12.5 | 15.2 | 10.5855 | 11.51 | 11.1 | 11.69 | 12.72 | 12.55 | 12.3 |  |  |  |  | 192
| 1 | 6 | 10.09 | 12.19 | 11.2 | 11.71 | 11.5 | 11.9 | 10.6 | 13.9 | 12.2 | 11.75 | 12.15 | 12.55 | 12.1 | 14.6 | 11.8656 | 11.8 | 11.1 | 11.73 | 12.6 | 12.21 | 12
| Mean |  | 10.465 | 11.965 | 11.25 | 11.76 | 11.25 | 12.1167 | 10.5833 | 13.8167 | 12.3083 | 11.8 | 12.5417 | 12.1417 | 12.6333 | 15.4667 | 11.1586 | 11.89 | 11.1 | 11.2867 | 12.825 | 12.335 | 11.9167
| Median |  | 10.2 | 11.995 | 11.25 | 11.735 | 11.25 | 12.05 | 10.6 | 13.8 | 12.2 | 11.725 | 12.525 | 12.1 | 12.5 | 15.15 | 11.2179 | 11.845 | 11.1 | 11.32 | 12.805 | 12.38 | 12
| Std Dev. |  | 0.53276 | 0.17003 | 0.23452 | 0.078994 | 0.22583 | 0.33116 | 0.098319 | 0.17224 | 0.28708 | 0.16125 | 0.29903 | 0.21775 | 0.85186 | 0.89368 | 0.64085 | 0.30318 | 0.089443 | 0.42852 | 0.23688 | 0.31923 | 0.31885
| Rel.Std.Dev. |  | 5.091% | 1.421% | 2.085% | 0.6717% | 2.007% | 2.733% | 0.929% | 1.247% | 2.332% | 1.366% | 2.384% | 1.793% | 6.743% | 5.778% | 5.743% | 2.55% | 0.8058% | 3.797% | 1.847% | 2.588% | 2.676%
| PDM3 |  | -11.63% | 1.033% | -5.005% | -0.6981% | -5.005% | 2.314% | -10.63% | 16.67% | 3.932% | -0.3604% | 5.902% | 2.525% | 6.676% | 30.6% | -5.777% | 0.3996% | -6.271% | -4.695% | 8.295% | 4.157% | 0.6248%
| Z-Score (Absolute) |  | 1.55 | 0.05 | 0.71 | 0.17 | 0.71 | 0.21 | 1.43 | 2.03 | 0.42 | 0.12 | 0.67 | 0.24 | 0.77 | 3.8 | 0.81 | 0.03 | 0.87 | 0.67 | 0.97 | 0.45 | 0
| Analytical results for V in OREAS 554 (Certified Value 19.3 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 25g | 01g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  | 2
| 1 | 1 | 20 | 22 | 50 | 26 | 20 | 18 | 25 | 20.4 | 20 | 20 | 20 | 19 | 20 | 19 | 5 | 17.0917 | 16 | 25 | 19 | 22 | 18 | 16.6 |  |  |  | 1
| 1 | 2 | 21 | 22 | 30 | 26 | 20 | 18 | 25 | 20.3 | 20 | 20 | 20 | 19 | 20 | 19 | 5 | 19.0355 | 16 | 25 | 20 | 22 | 19 | 16.6 |  |  |  | 0
| 1 | 3 | 20 | 22 | 20 | 25 | 20 | 17 | 25 | 20.7 | 20 | 20 | 20 | 19 | 20 | 19 | 5 | 16.4053 | 18 | 25 | 20 | 22 | 18 | 17.4 |  |  |  | 16
| 1 | 4 | 19 | 21 | 20 | 25 | 20 | 17 | 25 | 19.8 | 20 | 20 | 19 | 19 | 19 | 19 | 5 | 17.6861 | 17 | 25 | 20 | 22 | 18 | 17.1 |  |  |  | 19.2956
| 1 | 5 | 20 | 21 | 20 | 25 | 20 | 17 | 25 | 19.8 | 20 | 19 | 19 | 19 | 19 | 19 | 5 | 19.1014 | 18 | 25 | 19 | 22 | 19 | 16.8 |  |  |  | 193
| 1 | 6 | 19 | 22 | 20 | 25 | 20 | 18 | 24 | 20.2 | 21 | 20 | 19 | 19 | 20 | 19 | 5 | 19.3305 | 17 | 25 | 18 | 22 | 18 | 16.8
| Mean |  | 19.8333 | 21.6667 | 26.6667 | 25.3333 | 20 | 17.5 | 24.8333 | 20.2 | 20.1667 | 19.8333 | 19.5 | 19 | 19.6667 | 19 | 5 | 18.1084 | 17 | 25 | 19.3333 | 22 | 18.3333 | 16.8833
| Median |  | 20 | 22 | 20 | 25 | 20 | 17.5 | 25 | 20.25 | 20 | 20 | 19.5 | 19 | 20 | 19 | 5 | 18.3608 | 17 | 25 | 19.5 | 22 | 18 | 16.8
| Std Dev. |  | 0.75277 | 0.5164 | 12.1106 | 0.5164 | 0 | 0.54772 | 0.40825 | 0.35214 | 0.40825 | 0.40825 | 0.54772 | 0 | 0.5164 | 0 | 0 | 1.2208 | 0.89443 | 0 | 0.8165 | 0 | 0.5164 | 0.31252
| Rel.Std.Dev. |  | 3.795% | 2.383% | 45.41% | 2.038% | 0% | 3.13% | 1.644% | 1.743% | 2.024% | 2.058% | 2.809% | 0% | 2.626% | 0% | 0% | 6.742% | 5.261% | 0% | 4.223% | 0% | 2.817% | 1.851%
| PDM3 |  | 2.787% | 12.29% | 38.2% | 31.29% | 3.651% | -9.306% | 28.7% | 4.687% | 4.514% | 2.787% | 1.059% | -1.532% | 1.923% | -1.532% | -74.09% | -6.152% | -11.9% | 29.56% | 0.1956% | 14.02% | -4.987% | -12.5%
| Z-Score (Absolute) |  | 0.08 | 1.01 | NA | 2.87 | 0.17 | 1.1 | 2.61 | 0.27 | 0.25 | 0.08 | 0.08 | 0.34 | 0 | 0.34 | 7.42 | 0.79 | 1.35 | 2.7 | 0.17 | 1.18 | 0.67 | 1.41
Results from laboratory 4 were removed due to their 10 ppm reading resolution.
| Analytical results for W in OREAS 554 (Certified Value 0.1 ppm) | [26 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 23 | 24 | 28 | 29 | 30 |  |  |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 10 | 0.13 | < 0.5 | 0.1 | < 1 | < 0.5 | 0.1 | 0.3 | < 5 | 0.1 | 0.09 | 0.1 | 0.13 | 0.11 | 0.1 | 0.1 | 0.1 | 0.14 | 0.1 | 0.1 | 0.1 |  |  |  |  | 1
| 1 | 2 | < 10 | 0.13 | < 0.5 | 0.1 | < 1 | < 0.5 | 0.2 | 0.2 | < 5 | 0.11 | 0.09 | 0.1 | 0.11 | 0.11 | 0.1 | < 0.1 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 |  |  |  |  | 0
| 1 | 3 | < 10 | 0.13 | < 0.5 | 0.1 | < 1 | < 0.5 | 0.1 | 0.2 | < 5 | 0.11 | 0.08 | 0.11 | 0.1 | 0.11 | 0.2 | 0.1 | 0.1 | 0.11 | 0.11 | 0.1 | 0.1 |  |  |  |  | 16
| 1 | 4 | < 10 | 0.16 | < 0.5 | 0.1 | < 1 | < 0.5 | 0.1 | 0.2 | < 5 | 0.08 | 0.7 | 0.09 | 0.12 | 0.11 | 0.1 | 0.1 | 0.1 | 0.09 | 0.1 | 0.1 | 0.1 |  |  |  |  | 0.10471
| 1 | 5 | < 10 | 0.13 | < 0.5 | < 0.1 | < 1 | < 0.5 | 0.1 | 0.2 | < 5 | 0.1 | 0.11 | 0.11 | 0.12 | 0.09 | 0.1 | 0.1 | 0.1 | 0.13 | 0.11 | 0.1 | 0.1 |  |  |  |  | 194
| 1 | 6 | < 10 | 0.13 | < 0.5 | 0.1 | < 1 | < 0.5 | 0.1 | 0.2 | < 5 | 0.09 | 0.11 | 0.1 | 0.11 | 0.1 | 0.1 | 0.1 | 0.1 | 0.12 | 0.11 | 0.1 | 0.1
| Mean |  |  | 0.135 |  | 0.1 |  |  | 0.11667 | 0.21667 |  | 0.098333 | 0.19667 | 0.10167 | 0.115 | 0.105 | 0.11667 | 0.1 | 0.1 | 0.11667 | 0.105 | 0.1 | 0.1
| Median |  |  | 0.13 |  | 0.1 |  |  | 0.1 | 0.2 |  | 0.1 | 0.1 | 0.1 | 0.115 | 0.11 | 0.1 | 0.1 | 0.1 | 0.115 | 0.105 | 0.1 | 0.1
| Std Dev. |  |  | 0.012247 |  | 0 |  |  | 0.040825 | 0.040825 |  | 0.01169 | 0.24687 | 0.0075277 | 0.010488 | 0.0083666 | 0.040825 | 0 | 1.5202e-17 | 0.017512 | 0.0054772 | 1.5202e-17 | 1.5202e-17
| Rel.Std.Dev. |  |  | 9.072% |  | 0% |  |  | 34.99% | 18.84% |  | 11.89% | 125.5% | 7.404% | 9.12% | 7.968% | 34.99% | 0% | 1.52e-14% | 15.01% | 5.216% | 1.52e-14% | 1.52e-14%
| PDM3 |  |  | 28.93% |  | -4.497% |  |  | 11.42% | 106.9% |  | -6.088% | 87.82% | -2.905% | 9.829% | 0.2786% | 11.42% | -4.497% | -4.497% | 11.42% | 0.2786% | -4.497% | -4.497%
| Z-Score (Absolute) |  | NA | 0.67 | 4.9 | NA | 14.1 | 4.9 | NA | NA | 87.66 | 0.67 | 2.94 | 0.55 | 0.06 | 0.43 | NA | NA | NA | 0 | 0.43 | 0.61 | NA
Results from laboratories 1, 5, 10, 11, 18, 21, 23 & 30 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 554 (Certified Value 5.69 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 21 | 23 | 28 | 29 | 30 | [6 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 6.2 | 5.32 | 5.2 | 5.77 | 5.6 | 5.59 | 5.92 | 6 | 6.05 | 6.14 | 6.19 | 5.71 | 5.67 | 6.8 | 5.59 | 6.63 | 5.79 | 5.56 | 4.96 | [6 empty] | 1
| 1 | 2 | 6.6 | 5.47 | 5.25 | 5.84 | 5.3 | 5.58 | 5.76 | 6 | 6.01 | 5.95 | 6.42 | 5.64 | 5.59 | 6.8 | 5.47 | 6.42 | 5.86 | 5.71 | 4.88 | [6 empty] | 0
| 1 | 3 | 6 | 5.37 | 5.15 | 5.8 | 5.4 | 5.6 | 5.94 | 6 | 5.99 | 6.14 | 6.11 | 5.81 | 5.53 | 7.8 | 5.7 | 6.39 | 5.63 | 5.59 | 4.99 | [6 empty] | 16
| 1 | 4 | 5.6 | 5.2 | 5.1 | 5.44 | 5.4 | 5.65 | 5.71 | 6 | 5.75 | 6.07 | 6.2 | 5.68 | 5.57 | 7.2 | 5.74 | 6.43 | 5.75 | 5.71 | 5.02 | [6 empty] | 5.6926
| 1 | 5 | 5.8 | 5.19 | 5.15 | 5.89 | 5.3 | 5.6 | 5.76 | 6 | 5.75 | 5.84 | 6.27 | 5.37 | 5.58 | 6.9 | 5.61 | 6.21 | 5.7 | 5.71 | 4.89 | [6 empty] | 195
| 1 | 6 | 5.7 | 5.44 | 5.15 | 5.43 | 5.6 | 5.58 | 5.79 | 5 | 6.02 | 5.85 | 6.16 | 5.72 | 5.74 | 6.9 | 5.6 | 6.35 | 5.52 | 5.45 | 5.02
| Mean |  | 5.9833 | 5.3317 | 5.1667 | 5.695 | 5.4333 | 5.6 | 5.8133 | 5.8333 | 5.9283 | 5.9983 | 6.225 | 5.655 | 5.6133 | 7.0667 | 5.6183 | 6.405 | 5.7083 | 5.6217 | 4.96
| Median |  | 5.9 | 5.345 | 5.15 | 5.785 | 5.4 | 5.595 | 5.775 | 6 | 6 | 6.01 | 6.195 | 5.695 | 5.585 | 6.9 | 5.605 | 6.405 | 5.725 | 5.65 | 4.975
| Std Dev. |  | 0.37103 | 0.11822 | 0.05164 | 0.2054 | 0.13663 | 0.026077 | 0.094163 | 0.40825 | 0.13949 | 0.13761 | 0.10895 | 0.15057 | 0.077115 | 0.38816 | 0.094534 | 0.1362 | 0.1209 | 0.10741 | 0.06229
| Rel.Std.Dev. |  | 6.201% | 2.217% | 0.9995% | 3.607% | 2.515% | 0.4657% | 1.62% | 6.999% | 2.353% | 2.294% | 1.75% | 2.663% | 1.374% | 5.493% | 1.683% | 2.126% | 2.118% | 1.911% | 1.256%
| PDM3 |  | 5.107% | -6.341% | -9.239% | 0.04168% | -4.555% | -1.627% | 2.12% | 2.472% | 4.141% | 5.37% | 9.352% | -0.661% | -1.393% | 24.14% | -1.305% | 12.51% | 0.2759% | -1.247% | -12.87%
| Z-Score (Absolute) |  | 0.84 | 0.94 | 1.38 | 0.05 | 0.66 | 0.2 | 0.38 | NA | 0.69 | 0.88 | 1.5 | 0.05 | 0.17 | 3.79 | 0.15 | 1.99 | 0.09 | 0.15 | 1.95
Results from laboratory 12 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 554 (Certified Value 0.5 ppm) | [26 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 10: Lab; 11: Lab; 12: Lab | [6 empty] | 1
| No | No | 01: 02; 02: 04; 04: 06; 05: 10; 06: 11; 10: 21; 11: 29; 12: 30 | [6 empty] | ppm
|  |  | 01: AR*MS; 02: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*OES/MS; 10: AR*OES/MS; 11: AR*OES/MS; 12: AR*MS | [6 empty] | 2
|  |  | 01: 25g; 02: 01-05g; 04: 25g; 05: 0.5g; 06: 0.25g; 10: 0.25g; 11: 0.25g; 12: 0.5g | [6 empty] | 3
| 1 | 1 | 01: 0.494; 02: 0.495; 04: 0.55; 05: 0.5; 06: 0.7; 10: 0.5; 11: 0.5; 12: 0.51 | [6 empty] | 1
| 1 | 2 | 01: 0.506; 02: 0.49; 04: 0.5; 05: 0.5; 06: 0.7; 10: 0.5; 11: 0.5; 12: 0.48 | [6 empty] | 0
| 1 | 3 | 01: 0.528; 02: 0.5; 04: 0.55; 05: 0.5; 06: 0.8; 10: 0.5; 11: 0.5; 12: 0.51 | [6 empty] | 16
| 1 | 4 | 01: 0.501; 02: 0.485; 04: 0.55; 05: 0.5; 06: 0.7; 10: 0.5; 11: 0.5; 12: 0.49 | [6 empty] | 0.49776
| 1 | 5 | 01: 0.488; 02: 0.49; 04: 0.55; 05: 0.5; 06: 0.7; 10: 0.5; 11: 0.5; 12: 0.48 | [6 empty] | 196
| 1 | 6 | 01: 0.517; 02: 0.535; 04: 0.55; 05: 0.5; 06: 0.7; 10: 0.5; 11: 0.5; 12: 0.49
| Mean |  | 01: 0.50567; 02: 0.49917; 04: 0.54167; 05: 0.5; 06: 0.71667; 10: 0.5; 11: 0.5; 12: 0.49333
| Median |  | 01: 0.5035; 02: 0.4925; 04: 0.55; 05: 0.5; 06: 0.7; 10: 0.5; 11: 0.5; 12: 0.49
| Std Dev. |  | 01: 0.014814; 02: 0.01828; 04: 0.020412; 05: 0; 06: 0.040825; 10: 0; 11: 0; 12: 0.013663
| Rel.Std.Dev. |  | 01: 2.93%; 02: 3.662%; 04: 3.768%; 05: 0%; 06: 5.696%; 10: 0%; 11: 0%; 12: 2.769%
| PDM3 |  | 01: 1.589%; 02: 0.2835%; 04: 8.822%; 05: 0.4509%; 06: 43.98%; 10: 0.4509%; 11: 0.4509%; 12: -0.8884%
| Z-Score (Absolute) |  | 01: 1.18; 02: 0.17; 04: 8.64; 05: 0; 06: 44.95; 10: 0; 11: 0; 12: 1.38
| Analytical results for Zn in OREAS 554 (Certified Value 48.2 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 21 | 22 | 23 | 24 | 28 | 29 | 30 |  |  | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 47 | 49 | 47 | 51 | 50 | 55 | 49.7 | 45.1 | 41 | 49 | 48 | 49 | 49 | 49 | 38.8 | 79.2757 | 48 | 51.7571 | 62 | 51.6 | 43 | 50 | 50.9 |  |  | 1
| 1 | 2 | 48 | 52 | 47 | 50 | 45 | 53 | 48.6 | 44.2 | 42 | 48 | 47 | 48 | 48 | 47 | 40.8 | 61.7769 | 46 | 51.9759 | 64 | 52.2 | 43 | 50 | 49.1 |  |  | 0
| 1 | 3 | 48 | 50 | 45 | 51 | 50 | 55 | 49.7 | 45.1 | 40 | 48 | 48 | 49 | 47 | 47 | 38.7 | 53.8322 | 46 | 51.9677 | 63 | 52.1 | 44 | 50 | 51 |  |  | 16
| 1 | 4 | 45 | 48 | 45 | 50 | 50 | 54 | 50.5 | 44.4 | 41 | 47 | 46 | 47 | 47 | 46 | 39 | 52.3413 | 46 | 53.0966 | 63 | 52 | 44 | 50 | 50.5 |  |  | 48.2215
| 1 | 5 | 46 | 49 | 45 | 50 | 50 | 56 | 50.3 | 43.3 | 41 | 46 | 46 | 48 | 47 | 46 | 40 | 52.2864 | 48 | 52.2588 | 63 | 50.5 | 43 | 51 | 50.5 |  |  | 197
| 1 | 6 | 46 | 51 | 46 | 50 | 50 | 55 | 48.4 | 45.9 | 40 | 48 | 47 | 48 | 50 | 46 | 38.9 | 63.6516 | 49 | 52.5184 | 64 | 50.5 | 42 | 49 | 49.9
| Mean |  | 46.6667 | 49.8333 | 45.8333 | 50.3333 | 49.1667 | 54.6667 | 49.5333 | 44.6667 | 40.8333 | 47.6667 | 47 | 48.1667 | 48 | 46.8333 | 39.3667 | 60.5273 | 47.1667 | 52.2624 | 63.1667 | 51.4833 | 43.1667 | 50 | 50.3167
| Median |  | 46.5 | 49.5 | 45.5 | 50 | 50 | 55 | 49.7 | 44.75 | 41 | 48 | 47 | 48 | 47.5 | 46.5 | 38.95 | 57.8045 | 47 | 52.1174 | 63 | 51.8 | 43 | 50 | 50.5
| Std Dev. |  | 1.2111 | 1.472 | 0.98319 | 0.5164 | 2.0412 | 1.0328 | 0.8641 | 0.90037 | 0.75277 | 1.0328 | 0.89443 | 0.75277 | 1.2649 | 1.169 | 0.84538 | 10.417 | 1.3292 | 0.48699 | 0.75277 | 0.78846 | 0.75277 | 0.63246 | 0.7111
| Rel.Std.Dev. |  | 2.595% | 2.954% | 2.145% | 1.026% | 4.152% | 1.889% | 1.744% | 2.016% | 1.844% | 2.167% | 1.903% | 1.563% | 2.635% | 2.496% | 2.147% | 17.21% | 2.818% | 0.9318% | 1.192% | 1.531% | 1.744% | 1.265% | 1.413%
| PDM3 |  | -3.224% | 3.343% | -4.952% | 4.38% | 1.96% | 13.37% | 2.721% | -7.372% | -15.32% | -1.151% | -2.533% | -0.1136% | -0.4592% | -2.879% | -18.36% | 25.52% | -2.187% | 8.38% | 30.99% | 6.764% | -10.48% | 3.688% | 4.345%
| Z-Score (Absolute) |  | 0.47 | 0.52 | 0.73 | 0.68 | 0.31 | 2.04 | 0.43 | 1.1 | 2.3 | 0.16 | 0.37 | 0 | 0.05 | 0.42 | 2.76 | 3.88 | 0.31 | 1.28 | 4.7 | 1.04 | 1.57 | 0.57 | 0.67
| Analytical results for Zr in OREAS 554 (Certified Value 8.99 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 21 | 23 | 28 | 29 | 30 | [7 empty] | ppm
|  |  | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.25g | 25g | 01-05g | 01g | 25g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.25g | 0.5g | [7 empty] | 3
| 1 | 1 | 9.6 | 8.8 | 8 | 8.4 | 10 | 9.2 | 10.4 | 9 | 9.1 | 9 | 9.8 | 8.8 | 9.5 | 8.3 | 8.5 | 10.3 | 8.7 | 6.95 | [7 empty] | 1
| 1 | 2 | 9.5 | 9.2 | 7.5 | 8.4 | 11 | 9 | 10.2 | 9 | 9.1 | 8.8 | 9.8 | 8.5 | 9.1 | 8.7 | 8.8 | 10.4 | 8.9 | 6.94 | [7 empty] | 0
| 1 | 3 | 9.6 | 8.9 | 8 | 9.4 | 11 | 8.8 | 10.1 | 9 | 9.5 | 9 | 9.7 | 8.9 | 9.2 | 9.4 | 8.6 | 10 | 8.6 | 7.17 | [7 empty] | 16
| 1 | 4 | 8.9 | 8.5 | 7.5 | 8.8 | 11 | 9.1 | 10 | 9 | 8.8 | 9.2 | 9.7 | 8.6 | 9.1 | 8.7 | 8.4 | 10.5 | 8.7 | 7.08 | [7 empty] | 8.9917
| 1 | 5 | 9.1 | 8.7 | 7.5 | 9.1 | 10 | 9.2 | 9.9 | 9 | 9 | 9 | 10 | 8.1 | 8.9 | 9.2 | 8.5 | 10.1 | 9 | 6.91 | [7 empty] | 198
| 1 | 6 | 9.1 | 9 | 8 | 9.2 | 10 | 8.9 | 10.4 | 9 | 9.1 | 9 | 9.8 | 8.6 | 9 | 8.8 | 8.6 | 9.8 | 8.9 | 6.94
| Mean |  | 9.3 | 8.85 | 7.75 | 8.8833 | 10.5 | 9.0333 | 10.1667 | 9 | 9.1 | 9 | 9.8 | 8.5833 | 9.1333 | 8.85 | 8.5667 | 10.1833 | 8.8 | 6.9983
| Median |  | 9.3 | 8.85 | 7.75 | 8.95 | 10.5 | 9.05 | 10.15 | 9 | 9.1 | 9 | 9.8 | 8.6 | 9.1 | 8.75 | 8.55 | 10.2 | 8.8 | 6.945
| Std Dev. |  | 0.30332 | 0.2429 | 0.27386 | 0.42151 | 0.54772 | 0.1633 | 0.20656 | 0 | 0.22804 | 0.12649 | 0.10954 | 0.27869 | 0.20656 | 0.3937 | 0.13663 | 0.26394 | 0.15492 | 0.10304
| Rel.Std.Dev. |  | 3.261% | 2.745% | 3.534% | 4.745% | 5.216% | 1.808% | 2.032% | 0% | 2.506% | 1.405% | 1.118% | 3.247% | 2.262% | 4.449% | 1.595% | 2.592% | 1.76% | 1.472%
| PDM3 |  | 3.429% | -1.576% | -13.81% | -1.205% | 16.77% | 0.4634% | 13.07% | 0.09268% | 1.205% | 0.09268% | 8.99% | -4.541% | 1.576% | -1.576% | -4.727% | 13.25% | -2.132% | -22.17%
| Z-Score (Absolute) |  | 0.88 | 0.22 | 2.92 | 0.14 | NA | 0.22 | 3 | NA | 0.39 | 0.14 | 2.1 | 0.88 | 0.47 | 0.22 | 0.92 | 3.04 | 0.35 | 4.77
Results from laboratories 6 & 12 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C364: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C530: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C549: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C567: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C585: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C603: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C640: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C658: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C676: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C694: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C712: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C804: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C822: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1059: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1152: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1171: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1189: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1207: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.