## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 235b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 1.6338 | 0.052986 | 1.5279 | 1.7398 | 1.4749 | 1.7928 | 3.243% | 6.486% | 9.729% | 1.5521 | 1.7155
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.6042 | 0.0785 | 1.4472 | 1.7612 | 1.3687 | 1.8397 | 4.893% | 9.787% | 14.68% | 1.524 | 1.6845
Cyanide Leach
Au, ppm | 1.4298 | 0.074896 | 1.28 | 1.5796 | 1.2051 | 1.6544 | 5.238% | 10.48% | 15.71% | 1.3583 | 1.5012
PhotonAssay
Au, ppm | 1.6772 | 0.062266 | 1.5527 | 1.8018 | 1.4904 | 1.864 | 3.712% | 7.425% | 11.14% | 1.5934 | 1.7611
4-Acid Digestion
Ag, ppm | 0.13537 | 0.027629 | 0.080108 | 0.19063 | 0.052478 | 0.21825 | 20.41% | 40.82% | 61.23% | 0.1286 | 0.14214
Al, wt.% | 7.5667 | 0.22293 | 7.1208 | 8.0126 | 6.8979 | 8.2355 | 2.946% | 5.892% | 8.839% | 7.1884 | 7.945
As, ppm | 353.2216 | 16.1562 | 320.9091 | 385.534 | 304.7529 | 401.6902 | 4.574% | 9.148% | 13.72% | 335.5605 | 370.8826
Ba, ppm | 691.2618 | 21.3592 | 648.5434 | 733.9803 | 627.1841 | 755.3395 | 3.09% | 6.18% | 9.27% | 656.6987 | 725.8249
Be, ppm | 2.5425 | 0.21698 | 2.1085 | 2.9764 | 1.8915 | 3.1934 | 8.534% | 17.07% | 25.6% | 2.4153 | 2.6696
Bi, ppm | 0.34364 | 0.026466 | 0.29071 | 0.39657 | 0.26424 | 0.42304 | 7.702% | 15.4% | 23.1% | 0.32646 | 0.36082
Ca, wt.% | 0.85969 | 0.037516 | 0.78466 | 0.93472 | 0.74714 | 0.97224 | 4.364% | 8.728% | 13.09% | 0.8167 | 0.90267
Ce, ppm | 83.0499 | 3.3998 | 76.2504 | 89.8495 | 72.8506 | 93.2493 | 4.094% | 8.187% | 12.28% | 78.8974 | 87.2024
Co, ppm | 17.7651 | 0.50107 | 16.7629 | 18.7672 | 16.2618 | 19.2683 | 2.821% | 5.641% | 8.462% | 16.8768 | 18.6533
Cr, ppm | 128.2002 | 12.1906 | 103.8189 | 152.5815 | 91.6283 | 164.7721 | 9.509% | 19.02% | 28.53% | 121.7902 | 134.6102
Cs, ppm | 9.9752 | 0.49128 | 8.9927 | 10.9578 | 8.5014 | 11.4491 | 4.925% | 9.85% | 14.77% | 9.4765 | 10.474
Cu, ppm | 29.3305 | 1.469 | 26.3925 | 32.2685 | 24.9236 | 33.7375 | 5.008% | 10.02% | 15.03% | 27.864 | 30.797
Dy, ppm | 3.6851 | 0.69739 | 2.2903 | 5.0799 | 1.593 | 5.7773 | 18.92% | 37.85% | 56.77% | 3.5009 | 3.8694
Er, ppm | 2.0489 | 0.46395 | 1.121 | 2.9768 | 0.65702 | 3.4407 | 22.64% | 45.29% | 67.93% | 1.9464 | 2.1513
Eu, ppm | 1.2891 | 0.075194 | 1.1388 | 1.4395 | 1.0636 | 1.5147 | 5.833% | 11.67% | 17.5% | 1.2247 | 1.3536
Fe, wt.% | 4.1515 | 0.1239 | 3.9037 | 4.3993 | 3.7798 | 4.5232 | 2.984% | 5.969% | 8.953% | 3.944 | 4.3591
Ga, ppm | 19.8956 | 1.1808 | 17.5339 | 22.2572 | 16.3531 | 23.438 | 5.935% | 11.87% | 17.81% | 18.9008 | 20.8903
Gd, ppm | 5.2277 | 0.37415 | 4.4794 | 5.976 | 4.1053 | 6.3501 | 7.157% | 14.31% | 21.47% | 4.9663 | 5.4891
Hf, ppm | 3.9194 | 0.34692 | 3.2255 | 4.6132 | 2.8786 | 4.9601 | 8.851% | 17.7% | 26.55% | 3.7234 | 4.1153
Ho, ppm | 0.71004 | 0.16009 | 0.38986 | 1.0302 | 0.22978 | 1.1903 | 22.55% | 45.09% | 67.64% | 0.67454 | 0.74554
In, ppm | 0.069702 | 0.0051372 | 0.059427 | 0.079976 | 0.05429 | 0.085113 | 7.37% | 14.74% | 22.11% | 0.066216 | 0.073187
K, wt.% | 2.7593 | 0.071459 | 2.6164 | 2.9022 | 2.5449 | 2.9736 | 2.59% | 5.18% | 7.769% | 2.6213 | 2.8972
La, ppm | 40.6019 | 2.1369 | 36.3281 | 44.8758 | 34.1911 | 47.0127 | 5.263% | 10.53% | 15.79% | 38.5718 | 42.632
Li, ppm | 54.5974 | 3.487 | 47.6233 | 61.5714 | 44.1363 | 65.0585 | 6.387% | 12.77% | 19.16% | 51.8675 | 57.3273
Lu, ppm | 0.27208 | 0.056557 | 0.15896 | 0.38519 | 0.10241 | 0.44175 | 20.79% | 41.57% | 62.36% | 0.25847 | 0.28568
Mg, wt.% | 1.776 | 0.063845 | 1.6483 | 1.9037 | 1.5845 | 1.9675 | 3.595% | 7.19% | 10.78% | 1.6872 | 1.8648
Mn, wt.% | 0.040554 | 0.0016014 | 0.037351 | 0.043757 | 0.03575 | 0.045358 | 3.949% | 7.898% | 11.85% | 0.038526 | 0.042582
Mo, ppm | 0.9849 | 0.077573 | 0.82976 | 1.1401 | 0.75218 | 1.2176 | 7.876% | 15.75% | 23.63% | 0.93566 | 1.0341
Na, wt.% | 0.75371 | 0.017498 | 0.71872 | 0.78871 | 0.70122 | 0.80621 | 2.322% | 4.643% | 6.965% | 0.71603 | 0.7914
Nb, ppm | 15.8254 | 1.3064 | 13.2126 | 18.4382 | 11.9062 | 19.7446 | 8.255% | 16.51% | 24.77% | 15.0341 | 16.6167
Nd, ppm | 35.6873 | 1.3164 | 33.0545 | 38.32 | 31.7381 | 39.6364 | 3.689% | 7.377% | 11.07% | 33.9029 | 37.4716
Ni, ppm | 65.8957 | 2.8638 | 60.168 | 71.6234 | 57.3042 | 74.4872 | 4.346% | 8.692% | 13.04% | 62.6009 | 69.1905
P, wt.% | 0.075924 | 0.0030197 | 0.069884 | 0.081963 | 0.066865 | 0.084983 | 3.977% | 7.955% | 11.93% | 0.072128 | 0.07972
Pb, ppm | 19.0867 | 1.178 | 16.7308 | 21.4426 | 15.5528 | 22.6206 | 6.172% | 12.34% | 18.51% | 18.1324 | 20.041
Pr, ppm | 9.7137 | 0.53737 | 8.6389 | 10.7884 | 8.1015 | 11.3258 | 5.532% | 11.06% | 16.6% | 9.228 | 10.1993
Rb, ppm | 158.8882 | 10.9939 | 136.9004 | 180.876 | 125.9064 | 191.8699 | 6.919% | 13.84% | 20.76% | 150.9438 | 166.8326
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.1409 | 0.0075905 | 0.12572 | 0.15608 | 0.11813 | 0.16367 | 5.387% | 10.77% | 16.16% | 0.13385 | 0.14794
Sb, ppm | 327.4542 | 15.5037 | 296.4468 | 358.4615 | 280.9431 | 373.9652 | 4.735% | 9.469% | 14.2% | 311.0815 | 343.8269
Sc, ppm | 14.7174 | 0.94317 | 12.8311 | 16.6038 | 11.8879 | 17.547 | 6.409% | 12.82% | 19.23% | 13.9816 | 15.4533
Sm, ppm | 6.8023 | 0.45977 | 5.8828 | 7.7219 | 5.423 | 8.1817 | 6.759% | 13.52% | 20.28% | 6.4622 | 7.1425
Sn, ppm | 3.7912 | 0.25447 | 3.2822 | 4.3001 | 3.0278 | 4.5546 | 6.712% | 13.42% | 20.14% | 3.6016 | 3.9807
Sr, ppm | 149.4852 | 4.7721 | 139.9411 | 159.0294 | 135.169 | 163.8014 | 3.192% | 6.385% | 9.577% | 142.0109 | 156.9595
Ta, ppm | 1.1775 | 0.090054 | 0.99735 | 1.3576 | 0.9073 | 1.4476 | 7.648% | 15.3% | 22.94% | 1.1186 | 1.2363
Tb, ppm | 0.68857 | 0.047953 | 0.59267 | 0.78448 | 0.54471 | 0.83243 | 6.964% | 13.93% | 20.89% | 0.65414 | 0.723
Th, ppm | 14.8145 | 0.68403 | 13.4464 | 16.1826 | 12.7624 | 16.8666 | 4.617% | 9.235% | 13.85% | 14.0738 | 15.5552
Ti, wt.% | 0.48921 | 0.014069 | 0.46107 | 0.51735 | 0.447 | 0.53142 | 2.876% | 5.752% | 8.628% | 0.46475 | 0.51367
Tl, ppm | 0.85987 | 0.036752 | 0.78636 | 0.93337 | 0.74961 | 0.97012 | 4.274% | 8.548% | 12.82% | 0.81687 | 0.90286
Tm, ppm | 0.29508 | 0.080492 | 0.1341 | 0.45606 | 0.053606 | 0.53656 | 27.28% | 54.56% | 81.83% | 0.28033 | 0.30984
U, ppm | 2.8114 | 0.15889 | 2.4937 | 3.1292 | 2.3348 | 3.2881 | 5.651% | 11.3% | 16.95% | 2.6709 | 2.952
V, ppm | 110.9319 | 3.6509 | 103.63 | 118.2338 | 99.9791 | 121.8847 | 3.291% | 6.582% | 9.873% | 105.3853 | 116.4785
W, ppm | 2.3419 | 0.33557 | 1.6707 | 3.013 | 1.3351 | 3.3486 | 14.33% | 28.66% | 42.99% | 2.2248 | 2.4589
Y, ppm | 16.3526 | 2.6892 | 10.9742 | 21.731 | 8.2851 | 24.4202 | 16.44% | 32.89% | 49.33% | 15.535 | 17.1702
Yb, ppm | 1.8043 | 0.37739 | 1.0495 | 2.5591 | 0.67211 | 2.9365 | 20.92% | 41.83% | 62.75% | 1.7141 | 1.8945
Zn, ppm | 97.2547 | 3.5418 | 90.1711 | 104.3383 | 86.6293 | 107.8801 | 3.642% | 7.284% | 10.93% | 92.3919 | 102.1174
Zr, ppm | 139.6935 | 7.4061 | 124.8814 | 154.5056 | 117.4753 | 161.9117 | 5.302% | 10.6% | 15.9% | 132.7088 | 146.6782
Aqua Regia Digestion
Ag, ppm | 0.14502 | 0.010717 | 0.12358 | 0.16645 | 0.11287 | 0.17717 | 7.39% | 14.78% | 22.17% | 0.13776 | 0.15227
Al, wt.% | 3.0237 | 0.18644 | 2.6508 | 3.3966 | 2.4644 | 3.583 | 6.166% | 12.33% | 18.5% | 2.8725 | 3.1749
As, ppm | 352.156 | 16.7757 | 318.6047 | 385.7073 | 301.829 | 402.483 | 4.764% | 9.527% | 14.29% | 334.5482 | 369.7638
Ba, ppm | 96.0837 | 7.5798 | 80.9241 | 111.2432 | 73.3443 | 118.823 | 7.889% | 15.78% | 23.67% | 91.2795 | 100.8879
Be, ppm | 1.2774 | 0.16126 | 0.95486 | 1.5999 | 0.79359 | 1.7612 | 12.62% | 25.25% | 37.87% | 1.2135 | 1.3412
Bi, ppm | 0.33605 | 0.024418 | 0.28722 | 0.38489 | 0.2628 | 0.40931 | 7.266% | 14.53% | 21.8% | 0.31925 | 0.35286
Ca, wt.% | 0.33549 | 0.025437 | 0.28461 | 0.38636 | 0.25918 | 0.4118 | 7.582% | 15.16% | 22.75% | 0.31871 | 0.35226
Cd, ppm | 0.045903 | 0.015098 | 0.015708 | 0.076099 | 0.00060979 | 0.091197 | 32.89% | 65.78% | 98.67% | 0.043608 | 0.048199
Ce, ppm | 56.4326 | 6.6489 | 43.1347 | 69.7304 | 36.4858 | 76.3794 | 11.78% | 23.56% | 35.35% | 53.6109 | 59.2542
Co, ppm | 16.8668 | 0.87274 | 15.1213 | 18.6123 | 14.2486 | 19.485 | 5.174% | 10.35% | 15.52% | 16.0235 | 17.7102
Cr, ppm | 120.1468 | 5.2915 | 109.5638 | 130.7298 | 104.2723 | 136.0214 | 4.404% | 8.808% | 13.21% | 114.1395 | 126.1542
Cs, ppm | 7.4183 | 0.27804 | 6.8622 | 7.9743 | 6.5841 | 8.2524 | 3.748% | 7.496% | 11.24% | 7.0473 | 7.7892
Cu, ppm | 28.3834 | 1.6209 | 25.1417 | 31.6251 | 23.5208 | 33.246 | 5.711% | 11.42% | 17.13% | 26.9642 | 29.8026
Fe, wt.% | 3.6749 | 0.097653 | 3.4796 | 3.8702 | 3.3819 | 3.9679 | 2.657% | 5.315% | 7.972% | 3.4911 | 3.8586
Ga, ppm | 9.4568 | 0.47919 | 8.4985 | 10.4152 | 8.0193 | 10.8944 | 5.067% | 10.13% | 15.2% | 8.984 | 9.9297
Ge, ppm | 0.11531 | 0.020896 | 0.073514 | 0.1571 | 0.052618 | 0.178 | 18.12% | 36.24% | 54.37% | 0.10954 | 0.12107
Hf, ppm | 0.48776 | 0.078683 | 0.3304 | 0.64513 | 0.25172 | 0.72381 | 16.13% | 32.26% | 48.39% | 0.46338 | 0.51215
In, ppm | 0.037185 | 0.0028815 | 0.031422 | 0.042948 | 0.028541 | 0.04583 | 7.749% | 15.5% | 23.25% | 0.035326 | 0.039044
K, wt.% | 0.89959 | 0.049238 | 0.80111 | 0.99807 | 0.75188 | 1.0473 | 5.473% | 10.95% | 16.42% | 0.85461 | 0.94457
La, ppm | 29.1008 | 3.9134 | 21.274 | 36.9276 | 17.3606 | 40.8411 | 13.45% | 26.9% | 40.34% | 27.6458 | 30.5559
Li, ppm | 45.7841 | 3.0846 | 39.6149 | 51.9533 | 36.5303 | 55.0379 | 6.737% | 13.47% | 20.21% | 43.4949 | 48.0733
Lu, ppm | 0.12999 | 0.031941 | 0.066108 | 0.19387 | 0.034166 | 0.22581 | 24.57% | 49.14% | 73.72% | 0.12349 | 0.13649
Mg, wt.% | 1.4593 | 0.075266 | 1.3088 | 1.6098 | 1.2335 | 1.6851 | 5.158% | 10.32% | 15.47% | 1.3863 | 1.5323
Mn, wt.% | 0.030969 | 0.0017151 | 0.027539 | 0.0344 | 0.025824 | 0.036115 | 5.538% | 11.08% | 16.61% | 0.029421 | 0.032518
Mo, ppm | 0.93332 | 0.060203 | 0.81291 | 1.0537 | 0.75271 | 1.1139 | 6.45% | 12.9% | 19.35% | 0.88665 | 0.97998
Na, wt.% | 0.15503 | 0.013429 | 0.12817 | 0.18189 | 0.11474 | 0.19532 | 8.662% | 17.32% | 25.99% | 0.14728 | 0.16278
Nb, ppm | 0.52697 | 0.15017 | 0.22664 | 0.82731 | 0.076471 | 0.97748 | 28.5% | 56.99% | 85.49% | 0.50062 | 0.55332
Ni, ppm | 61.3819 | 4.0697 | 53.2424 | 69.5214 | 49.1726 | 73.5911 | 6.63% | 13.26% | 19.89% | 58.3128 | 64.451
P, wt.% | 0.067931 | 0.0017952 | 0.064341 | 0.071521 | 0.062545 | 0.073316 | 2.643% | 5.285% | 7.928% | 0.064534 | 0.071327
Pb, ppm | 9.4697 | 0.68637 | 8.097 | 10.8425 | 7.4107 | 11.5288 | 7.248% | 14.5% | 21.74% | 8.9963 | 9.9432
Rb, ppm | 88.1365 | 3.4736 | 81.1892 | 95.0837 | 77.7156 | 98.5574 | 3.941% | 7.882% | 11.82% | 83.7297 | 92.5433
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.14348 | 0.01267 | 0.11814 | 0.16882 | 0.10547 | 0.18149 | 8.83% | 17.66% | 26.49% | 0.13631 | 0.15066
Sb, ppm | 238.2006 | 45.8376 | 146.5254 | 329.8757 | 100.6879 | 375.7133 | 19.24% | 38.49% | 57.73% | 226.2906 | 250.1106
Sc, ppm | 7.8245 | 0.64687 | 6.5308 | 9.1182 | 5.8839 | 9.7651 | 8.267% | 16.53% | 24.8% | 7.4333 | 8.2157
Sn, ppm | 1.7558 | 0.11274 | 1.5303 | 1.9813 | 1.4176 | 2.0941 | 6.421% | 12.84% | 19.26% | 1.668 | 1.8436
Sr, ppm | 40.5041 | 2.1755 | 36.153 | 44.8552 | 33.9775 | 47.0307 | 5.371% | 10.74% | 16.11% | 38.4789 | 42.5293
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 12.8583 | 0.53617 | 11.786 | 13.9307 | 11.2498 | 14.4668 | 4.17% | 8.34% | 12.51% | 12.2154 | 13.5013
Ti, wt.% | 0.19283 | 0.01345 | 0.16593 | 0.21973 | 0.15248 | 0.23318 | 6.975% | 13.95% | 20.93% | 0.18319 | 0.20247
Tl, ppm | 0.53182 | 0.048474 | 0.43487 | 0.62877 | 0.3864 | 0.67724 | 9.115% | 18.23% | 27.34% | 0.50523 | 0.55841
U, ppm | 1.5164 | 0.10185 | 1.3127 | 1.7201 | 1.2109 | 1.8219 | 6.716% | 13.43% | 20.15% | 1.4406 | 1.5922
V, ppm | 80.8636 | 3.6952 | 73.4733 | 88.2539 | 69.7781 | 91.9491 | 4.57% | 9.139% | 13.71% | 76.8204 | 84.9068
W, ppm | 0.4719 | 0.062026 | 0.34784 | 0.59595 | 0.28582 | 0.65798 | 13.14% | 26.29% | 39.43% | 0.4483 | 0.49549
Y, ppm | 9.5672 | 0.7413 | 8.0846 | 11.0498 | 7.3433 | 11.7911 | 7.748% | 15.5% | 23.24% | 9.0888 | 10.0455
Yb, ppm | 0.95465 | 0.13495 | 0.68474 | 1.2245 | 0.54979 | 1.3595 | 14.14% | 28.27% | 42.41% | 0.90691 | 1.0024
Zn, ppm | 89.4789 | 3.2107 | 83.0575 | 95.9004 | 79.8468 | 99.1111 | 3.588% | 7.176% | 10.76% | 85.005 | 93.9529
Zr, ppm | 20.8776 | 3.6675 | 13.5427 | 28.2125 | 9.8752 | 31.88 | 17.57% | 35.13% | 52.7% | 19.8337 | 21.9215
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 235b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 1.6338 | 1.614 | 1.6537 | 1.6245 | 1.6432
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.6042 | 1.5701 | 1.6384 | 1.5942 | 1.6143
Cyanide Leach
Au, Gold (ppm) | 1.4298 | 1.3906 | 1.4689 | 1.4216 | 1.4379
PhotonAssay
Au, Gold (ppm) | 1.6772 | 1.6465 | 1.708 | 1.6744 | 1.68
4-Acid Digestion
Ag, Silver (ppm) | 0.13537 | 0.10382 | 0.16691 | 0.12208 | 0.14866
Al, Aluminium (wt.%) | 7.5667 | 7.3848 | 7.7486 | 7.3983 | 7.7351
As, Arsenic (ppm) | 353.2216 | 340.2683 | 366.1748 | 341.7091 | 364.734
Ba, Barium (ppm) | 691.2618 | 676.6349 | 705.8888 | 678.0221 | 704.5016
Be, Beryllium (ppm) | 2.5425 | 2.3569 | 2.728 | 2.4329 | 2.652
Bi, Bismuth (ppm) | 0.34364 | 0.30789 | 0.37938 | 0.30406 | 0.38322
Ca, Calcium (wt.%) | 0.85969 | 0.83619 | 0.88318 | 0.83707 | 0.88231
Ce, Cerium (ppm) | 83.0499 | 80.566 | 85.5338 | 80.4988 | 85.6011
Co, Cobalt (ppm) | 17.7651 | 17.0379 | 18.4922 | 17.2946 | 18.2355
Cr, Chromium (ppm) | 128.2002 | 120.5741 | 135.8262 | 124.5885 | 131.8118
Cs, Caesium (ppm) | 9.9752 | 9.6011 | 10.3493 | 9.6322 | 10.3183
Cu, Copper (ppm) | 29.3305 | 28.2526 | 30.4084 | 28.3144 | 30.3466
Dy, Dysprosium (ppm) | 3.6851 | 3.1405 | 4.2298 | 3.4753 | 3.8949
Er, Erbium (ppm) | 2.0489 | 1.7103 | 2.3875 | 1.9228 | 2.175
Eu, Europium (ppm) | 1.2891 | 1.2022 | 1.3761 | 1.2395 | 1.3388
Fe, Iron (wt.%) | 4.1515 | 4.0539 | 4.2492 | 4.0919 | 4.2112
Ga, Gallium (ppm) | 19.8956 | 18.8453 | 20.9459 | 19.3264 | 20.4647
Gd, Gadolinium (ppm) | 5.2277 | 4.9015 | 5.5539 | 4.9826 | 5.4728
Hf, Hafnium (ppm) | 3.9194 | 3.6293 | 4.2094 | 3.7666 | 4.0721
Ho, Holmium (ppm) | 0.71004 | 0.54535 | 0.87472 | 0.668 | 0.75208
In, Indium (ppm) | 0.069702 | 0.06127 | 0.078133 | 0.065124 | 0.074279
K, Potassium (wt.%) | 2.7593 | 2.689 | 2.8295 | 2.7067 | 2.8118
La, Lanthanum (ppm) | 40.6019 | 38.921 | 42.2828 | 39.3528 | 41.8511
Li, Lithium (ppm) | 54.5974 | 52.5824 | 56.6123 | 53.0706 | 56.1242
Lu, Lutetium (ppm) | 0.27208 | 0.20501 | 0.33914 | 0.25158 | 0.29257
Mg, Magnesium (wt.%) | 1.776 | 1.7365 | 1.8155 | 1.742 | 1.8101
Mn, Manganese (wt.%) | 0.040554 | 0.0396 | 0.041508 | 0.039781 | 0.041327
Mo, Molybdenum (ppm) | 0.9849 | 0.90285 | 1.067 | 0.89778 | 1.072
Na, Sodium (wt.%) | 0.75371 | 0.73454 | 0.77289 | 0.73597 | 0.77146
Nb, Niobium (ppm) | 15.8254 | 14.9379 | 16.713 | 15.3174 | 16.3334
Nd, Neodymium (ppm) | 35.6873 | 34.3969 | 36.9776 | 34.5002 | 36.8743
Ni, Nickel (ppm) | 65.8957 | 64.4181 | 67.3733 | 64.6554 | 67.136
P, Phosphorus (wt.%) | 0.075924 | 0.073973 | 0.077875 | 0.074758 | 0.077089
Pb, Lead (ppm) | 19.0867 | 17.9293 | 20.2441 | 17.8294 | 20.344
Pr, Praseodymium (ppm) | 9.7137 | 9.2681 | 10.1592 | 9.3709 | 10.0564
Rb, Rubidium (ppm) | 158.8882 | 150.5431 | 167.2333 | 152.6864 | 165.09
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.1409 | 0.1332 | 0.14859 | 0.13295 | 0.14885
Sb, Antimony (ppm) | 327.4542 | 313.6966 | 341.2117 | 319.3769 | 335.5315
Sc, Scandium (ppm) | 14.7174 | 13.8933 | 15.5416 | 14.3304 | 15.1045
Sm, Samarium (ppm) | 6.8023 | 6.3621 | 7.2426 | 6.46 | 7.1447
Sn, Tin (ppm) | 3.7912 | 3.5745 | 4.0078 | 3.6266 | 3.9558
Sr, Strontium (ppm) | 149.4852 | 146.0898 | 152.8806 | 146.3742 | 152.5962
Ta, Tantalum (ppm) | 1.1775 | 1.1123 | 1.2426 | 1.1325 | 1.2224
Tb, Terbium (ppm) | 0.68857 | 0.63967 | 0.73747 | 0.65026 | 0.72688
Th, Thorium (ppm) | 14.8145 | 14.2888 | 15.3402 | 14.3632 | 15.2658
Ti, Titanium (wt.%) | 0.48921 | 0.47629 | 0.50213 | 0.47806 | 0.50036
Tl, Thallium (ppm) | 0.85987 | 0.82223 | 0.8975 | 0.83482 | 0.88491
Tm, Thulium (ppm) | 0.29508 | 0.22045 | 0.36972 | 0.27683 | 0.31333
U, Uranium (ppm) | 2.8114 | 2.6898 | 2.9331 | 2.7277 | 2.8952
V, Vanadium (ppm) | 110.9319 | 108.0429 | 113.821 | 108.5907 | 113.2731
W, Tungsten (ppm) | 2.3419 | 2.088 | 2.5957 | 2.2022 | 2.4815
Y, Yttrium (ppm) | 16.3526 | 14.9946 | 17.7107 | 15.6946 | 17.0106
Yb, Ytterbium (ppm) | 1.8043 | 1.5437 | 2.0649 | 1.6864 | 1.9222
Zn, Zinc (ppm) | 97.2547 | 94.234 | 100.2754 | 95.1388 | 99.3705
Zr, Zirconium (ppm) | 139.6935 | 134.1543 | 145.2327 | 135.5696 | 143.8173
Aqua Regia Digestion
Ag, Silver (ppm) | 0.14502 | 0.12835 | 0.16168 | IND | IND
Al, Aluminium (wt.%) | 3.0237 | 2.9095 | 3.1379 | 2.9703 | 3.0771
As, Arsenic (ppm) | 352.156 | 340.7122 | 363.5998 | 341.3805 | 362.9315
Ba, Barium (ppm) | 96.0837 | 91.3131 | 100.8542 | 93.5024 | 98.665
Be, Beryllium (ppm) | 1.2774 | 1.1683 | 1.3865 | 1.2174 | 1.3374
Bi, Bismuth (ppm) | 0.33605 | 0.30919 | 0.36292 | 0.29498 | 0.37713
Ca, Calcium (wt.%) | 0.33549 | 0.32048 | 0.3505 | 0.32757 | 0.34341
Cd, Cadmium (ppm) | 0.045903 | 0.033981 | 0.057826 | 0.04205 | 0.049756
Ce, Cerium (ppm) | 56.4326 | 52.2093 | 60.6559 | 54.4171 | 58.448
Co, Cobalt (ppm) | 16.8668 | 16.2629 | 17.4708 | 16.3234 | 17.4103
Cr, Chromium (ppm) | 120.1468 | 116.6357 | 123.658 | 117.8783 | 122.4154
Cs, Caesium (ppm) | 7.4183 | 7.1072 | 7.7293 | 7.1655 | 7.671
Cu, Copper (ppm) | 28.3834 | 27.2146 | 29.5522 | 27.2168 | 29.55
Fe, Iron (wt.%) | 3.6749 | 3.6088 | 3.741 | 3.6266 | 3.7231
Ga, Gallium (ppm) | 9.4568 | 9.07 | 9.8437 | 9.2 | 9.7137
Ge, Germanium (ppm) | 0.11531 | 0.085221 | 0.14539 | IND | IND
Hf, Hafnium (ppm) | 0.48776 | 0.42666 | 0.54887 | 0.4626 | 0.51293
In, Indium (ppm) | 0.037185 | 0.034796 | 0.039575 | 0.03353 | 0.04084
K, Potassium (wt.%) | 0.89959 | 0.87251 | 0.92667 | 0.88506 | 0.91412
La, Lanthanum (ppm) | 29.1008 | 26.8655 | 31.3362 | 28.1137 | 30.088
Li, Lithium (ppm) | 45.7841 | 43.518 | 48.0502 | 44.5692 | 46.999
Lu, Lutetium (ppm) | 0.12999 | 0.084906 | 0.17507 | IND | IND
Mg, Magnesium (wt.%) | 1.4593 | 1.4163 | 1.5023 | 1.435 | 1.4836
Mn, Manganese (wt.%) | 0.030969 | 0.030016 | 0.031923 | 0.030462 | 0.031477
Mo, Molybdenum (ppm) | 0.93332 | 0.88477 | 0.98187 | 0.87967 | 0.98696
Na, Sodium (wt.%) | 0.15503 | 0.14513 | 0.16494 | 0.14947 | 0.1606
Nb, Niobium (ppm) | 0.52697 | 0.3978 | 0.65614 | 0.47891 | 0.57503
Ni, Nickel (ppm) | 61.3819 | 59.0621 | 63.7016 | 59.3879 | 63.3759
P, Phosphorus (wt.%) | 0.067931 | 0.066623 | 0.069239 | 0.066885 | 0.068977
Pb, Lead (ppm) | 9.4697 | 8.8243 | 10.1152 | 8.8668 | 10.0727
Rb, Rubidium (ppm) | 88.1365 | 84.8014 | 91.4716 | 85.9279 | 90.345
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.14348 | 0.13448 | 0.15248 | 0.13637 | 0.15059
Sb, Antimony (ppm) | 238.2006 | 215.6549 | 260.7463 | 231.4028 | 244.9984
Sc, Scandium (ppm) | 7.8245 | 7.3891 | 8.2599 | 7.6346 | 8.0144
Sn, Tin (ppm) | 1.7558 | 1.6506 | 1.8611 | 1.6677 | 1.844
Sr, Strontium (ppm) | 40.5041 | 38.2303 | 42.7779 | 39.4662 | 41.5419
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Th, Thorium (ppm) | 12.8583 | 12.2281 | 13.4885 | 12.4605 | 13.2562
Ti, Titanium (wt.%) | 0.19283 | 0.18471 | 0.20095 | 0.18839 | 0.19728
Tl, Thallium (ppm) | 0.53182 | 0.49442 | 0.56922 | 0.5133 | 0.55034
U, Uranium (ppm) | 1.5164 | 1.4362 | 1.5966 | 1.4639 | 1.5689
V, Vanadium (ppm) | 80.8636 | 78.7155 | 83.0118 | 79.5633 | 82.1639
W, Tungsten (ppm) | 0.4719 | 0.41331 | 0.53048 | 0.40513 | 0.53866
Y, Yttrium (ppm) | 9.5672 | 9.0452 | 10.0891 | 9.322 | 9.8123
Yb, Ytterbium (ppm) | 0.95465 | 0.77952 | 1.1298 | 0.85304 | 1.0563
Zn, Zinc (ppm) | 89.4789 | 86.8844 | 92.0735 | 87.4152 | 91.5426
Zr, Zirconium (ppm) | 20.8776 | 18.56 | 23.1952 | 20.274 | 21.4812
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 235b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 0.71898 | Pt | ppb | 0.52246
4-Acid Digestion
Au | ppm | 1.7112 | Hg | ppm | 0.025383 | Rh | ppm | 0.0025
B | ppm | 30.5583 | Ir | ppm | 0.0023667 | Ru | ppm | 0.0023833
Cd | ppm | 0.088371 | Pd | ppb | 58.25 | Se | ppm | 0.79305
Ge | ppm | 0.20666 | Pt | ppb | 7.7167 | Te | ppm | 0.055267
Aqua Regia Digestion
B | ppm | 15.9372 | Ir | ppm | < 0.0016 | Se | ppm | 0.23465
Dy | ppm | 2.1437 | Nd | ppm | 24.0205 | Si | wt.% | 0.47685
Er | ppm | 0.97779 | Pd | ppb | 37.25 | Sm | ppm | 4.4205
Eu | ppm | 0.62291 | Pr | ppm | 6.4845 | Tb | ppm | 0.50544
Gd | ppm | 4.1458 | Pt | ppb | 2.3 | Te | ppm | 0.028952
Hg | ppm | 0.012911 | Rh | ppm | 0.00083333 | Tm | ppm | 0.12344
Ho | ppm | 0.36191 | Ru | ppm | < 0.0065
Borate Fusion XRF
Al2O3 | wt.% | 15.025 | MgO | wt.% | 3.01 | SiO2 | wt.% | 66.805
CaO | wt.% | 1.21 | MnO | wt.% | 0.05 | SO3 | wt.% | 0.3565
Fe2O3 | wt.% | 6.06 | Na2O | wt.% | 1.04 | TiO2 | wt.% | 0.865
K2O | wt.% | 3.44 | P2O5 | wt.% | 0.174
Thermogravimetry
LOI1000 | wt.% | 2.145
Infrared Combustion
C | wt.% | 0.065 | S | wt.% | 0.13417
Laser Ablation ICP-MS
Ag | ppm | 0.2 | Hf | ppm | 6.375 | Sm | ppm | 7.335
As | ppm | 363.5 | Ho | ppm | 1.185 | Sn | ppm | 3.8
Ba | ppm | 692.5 | In | ppm | 0.0625 | Sr | ppm | 148
Be | ppm | 3 | La | ppm | 42.3 | Ta | ppm | 1.335
Bi | ppm | 0.37 | Lu | ppm | 0.49 | Tb | ppm | 0.99
Cd | ppm | 0.2 | Mn | wt.% | 0.044 | Te | ppm | < 0.2
Ce | ppm | 82.7 | Mo | ppm | 1.2 | Th | ppm | 15.65
Co | ppm | 18.75 | Nb | ppm | 17.7 | Ti | wt.% | 0.519
Cr | ppm | 151.5 | Nd | ppm | 37.95 | Tl | ppm | 0.4
Cs | ppm | 9.94 | Ni | ppm | 71 | Tm | ppm | 0.525
Cu | ppm | 31 | Pb | ppm | 21 | U | ppm | 3.23
Dy | ppm | 5.715 | Pr | ppm | 10.15 | V | ppm | 115.5
Er | ppm | 3.375 | Rb | ppm | 164.5 | W | ppm | 3
Eu | ppm | 1.41 | Re | ppm | < 0.01 | Y | ppm | 31.4
Ga | ppm | 19.6 | Sb | ppm | 345.5 | Yb | ppm | 3.37
Gd | ppm | 6.205 | Sc | ppm | 15.35 | Zn | ppm | 100
Ge | ppm | 1.55 | Se | ppm | < 5 | Zr | ppm | 232.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 235b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | Raw ~350g sample packed into 300cc jar analysed by PhotonAssay (high-power X-ray activation)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 235b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Hermosillo, Sonora, Mexico
Carsurin, Cikarang, West Java, Indonesia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Perth, WA, Australia
MinAnalytical Services, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 235b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
2 | 1 | ANSLu | 1 | 15 | 15 | 219494 | 0.082413 | 1.7051 | 1.7816
2 | 1 | ANSLu | 1 | 20 | 20 | 219495 | 0.086034 | 1.7944 | 1.7864
2 | 1 | ANSLu | 1 | 8 | 8 | 219496 | 0.083444 | 1.838 | 1.7887
2 | 1 | ANSLu | 1 | 16 | 16 | 219497 | 0.086093 | 1.6971 | 1.7812
2 | 1 | ANSLu | 1 | 11 | 11 | 219498 | 0.085303 | 1.8997 | 1.792
2 | 1 | ANSLu | 1 | 4 | 4 | 219499 | 0.08427 | 1.7099 | 1.7819
2 | 1 | ANSLu | 1 | 1 | 1 | 219500 | 0.087084 | 1.8006 | 1.7867
2 | 1 | ANSLu | 1 | 2 | 2 | 219501 | 0.084311 | 1.822 | 1.7878
2 | 1 | ANSLu | 1 | 10 | 10 | 219502 | 0.083342 | 1.8058 | 1.787
2 | 1 | ANSLu | 1 | 12 | 12 | 219503 | 0.083555 | 1.8108 | 1.7872
2 | 1 | ANSLu | 1 | 7 | 7 | 219504 | 0.086163 | 1.7649 | 1.7848
2 | 1 | ANSLu | 1 | 3 | 3 | 219505 | 0.084553 | 1.7607 | 1.7846
2 | 1 | ANSLu | 1 | 18 | 18 | 219506 | 0.086857 | 1.7545 | 1.7843
2 | 1 | ANSLu | 1 | 14 | 14 | 219507 | 0.087591 | 1.7992 | 1.7866
2 | 1 | ANSLu | 1 | 19 | 19 | 219508 | 0.083569 | 1.8733 | 1.7906
2 | 1 | ANSLu | 1 | 6 | 6 | 219509 | 0.084156 | 1.7824 | 1.7857
2 | 1 | ANSLu | 1 | 13 | 13 | 219510 | 0.085279 | 1.686 | 1.7806
2 | 1 | ANSLu | 1 | 9 | 9 | 219511 | 0.086194 | 1.8844 | 1.7912
2 | 1 | ANSLu | 1 | 5 | 5 | 219512 | 0.082737 | 1.7833 | 1.7858
2 | 1 | ANSLu | 1 | 17 | 17 | 219513 | 0.085848 | 1.7463 | 1.7838
Mean | [6 empty] | 0.08494 | 1.7859 | 1.7859
Median | [7 empty] | 1.7889 | 1.7861
SD | [7 empty] | 0.060516 | 0.0032201
RSD | [7 empty] | 3.388% | 0.1803%
RSD Upscale Factor | [7 empty] | 0.05321
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 235b (Certified Value 1.63 ppm) | [32 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*MS | 2
|  |  | 0.085g | 25g | 50g | 50g | 40g | 30g | 40g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 25g | 25g | 3
| 1 | 1 | 1.8006 | 1.686 | 1.71 | 1.509 | 1.6 | 1.6 | 1.66 | 1.61 | 1.64 | 1.63 | 1.663 | 1.655 | 1.605 | 1.665 | 1.67 | 1.678 | 1.5767 | 1.686 | 1.73 | 1.6707 | 1.67 | 1.56 | 1.622 | 1.48 | 1.609 | 1.58 | 1.608 | 1.671 | 1.5632 | 1.68 | 1.5783 | 1
| 1 | 2 | 1.822 | 1.727 | 1.7 | 1.515 | 1.63 | 1.58 | 1.64 | 1.58 | 1.64 | 1.67 | 1.631 | 1.66 | 1.625 | 1.645 | 1.695 | 1.681 | 1.5852 | 1.736 | 1.85 | 1.572 | 1.72 | 1.54 | 1.646 | 1.51 | 1.607 | 1.56 | 1.66 | 1.678 | 1.5592 | 1.66 | 1.5479 | 3
| 1 | 3 | 1.7607 | 1.7 | 1.75 | 1.441 | 1.62 | 1.61 | 1.66 | 1.6 | 1.64 | 1.63 | 1.597 | 1.675 | 1.65 | 1.61 | 1.66 | 1.678 | 1.6815 | 1.665 | 1.78 | 1.6673 | 1.65 | 1.58 | 1.629 | 1.51 | 1.618 | 1.57 | 1.74 | 1.679 | 1.584 | 1.71 | 1.5515 | 16
| 1 | 4 | 1.7099 | 1.652 | 1.68 | 1.502 | 1.64 | 1.58 | 1.66 | 1.61 | 1.66 | 1.62 | 1.57 | 1.665 | 1.605 | 1.665 | 1.645 | 1.662 | 1.5975 | 1.69 | 1.78 | 1.7031 | 1.65 | 1.61 | 1.633 | 1.49 | 1.59 | 1.62 | 1.68 | 1.71 | 1.5648 | 1.68 | 1.5922 | 1.6338
| 1 | 5 | 1.7833 | 1.803 | 1.69 | 1.437 | 1.65 | 1.63 | 1.67 | 1.65 | 1.63 | 1.62 | 1.654 | 1.67 | 1.615 | 1.675 | 1.68 | 1.656 | 1.565 | 1.698 | 1.84 | 1.5351 | 1.75 | 1.59 | 1.61 | 1.54 | 1.598 | 1.57 | 1.69 | 1.694 | 1.5744 | 1.69 | 1.5791 | 7
| 1 | 6 | 1.7824 | 1.725 | 1.63 | 1.484 | 1.62 | 1.6 | 1.68 | 1.67 | 1.62 | 1.65 | 1.622 | 1.67 | 1.64 | 1.66 | 1.665 | 1.656 | 1.6009 | 1.66 | 1.7 | 1.63 | 1.7 | 1.54 | 1.635 | 1.5 | 1.615 | 1.66 | 1.68 | 1.662 | 1.632 | 1.68 | 1.5446
|  | 7 | 2.00: 1.7649
|  | 8 | 2.00: 1.838
|  | 9 | 2.00: 1.8844
|  | 10 | 2.00: 1.8058
|  | 11 | 2.00: 1.8997
|  | 12 | 2.00: 1.8108
|  | 13 | 2.00: 1.686
|  | 14 | 2.00: 1.7992
|  | 15 | 2.00: 1.7051
|  | 16 | 2.00: 1.6971
|  | 17 | 2.00: 1.7463
|  | 18 | 2.00: 1.7545
|  | 19 | 2.00: 1.8733
|  | 20 | 2.00: 1.7944
| Mean |  | 1.7859 | 1.7155 | 1.6933 | 1.4813 | 1.6267 | 1.6 | 1.6617 | 1.62 | 1.6383 | 1.6367 | 1.6228 | 1.6658 | 1.6233 | 1.6533 | 1.6692 | 1.6685 | 1.6011 | 1.6892 | 1.78 | 1.6297 | 1.69 | 1.57 | 1.6292 | 1.505 | 1.6062 | 1.5933 | 1.6763 | 1.6823 | 1.5796 | 1.6833 | 1.5656
| Median |  | 1.7889 | 1.7125 | 1.695 | 1.493 | 1.625 | 1.6 | 1.66 | 1.61 | 1.64 | 1.63 | 1.6265 | 1.6675 | 1.62 | 1.6625 | 1.6675 | 1.67 | 1.5913 | 1.688 | 1.78 | 1.6487 | 1.685 | 1.57 | 1.631 | 1.505 | 1.608 | 1.575 | 1.68 | 1.6785 | 1.5696 | 1.68 | 1.5649
| Std Dev. |  | 0.060516 | 0.051032 | 0.039328 | 0.034425 | 0.017512 | 0.018974 | 0.013292 | 0.033466 | 0.013292 | 0.019664 | 0.03496 | 0.0073598 | 0.018619 | 0.023381 | 0.017151 | 0.01176 | 0.041563 | 0.027265 | 0.058992 | 0.064439 | 0.040497 | 0.028284 | 0.012254 | 0.020736 | 0.010534 | 0.038816 | 0.042903 | 0.017166 | 0.027189 | 0.01633 | 0.02003
| Rel.Std.Dev. |  | 3.388% | 2.975% | 2.323% | 2.324% | 1.077% | 1.186% | 0.7999% | 2.066% | 0.8113% | 1.201% | 2.154% | 0.4418% | 1.147% | 1.414% | 1.028% | 0.7048% | 2.596% | 1.614% | 3.314% | 3.954% | 2.396% | 1.802% | 0.7522% | 1.378% | 0.6559% | 2.436% | 2.559% | 1.02% | 1.721% | 0.9701% | 1.279%
| PDM3 |  | 9.309% | 4.999% | 3.642% | -9.334% | -0.4385% | -2.071% | 1.704% | -0.8465% | 0.2756% | 0.1736% | -0.6731% | 1.959% | -0.6425% | 1.194% | 2.163% | 2.122% | -2.002% | 3.387% | 8.946% | -0.2535% | 3.438% | -3.907% | -0.2855% | -7.885% | -1.693% | -2.479% | 2.601% | 2.969% | -3.319% | 3.03% | -4.177%
| Z-Score (Absolute) |  | NA | 1.56 | 1.14 | 2.87 | 0.12 | 0.63 | 0.54 | 0.25 | 0.1 | 0.07 | 0.2 | 0.62 | 0.19 | 0.38 | 0.68 | 0.67 | 0.61 | 1.06 | 2.78 | 0.07 | 1.07 | 1.19 | 0.08 | 2.42 | 0.51 | 0.75 | 0.82 | 0.93 | 1.01 | 0.95 | 1.28
| Analytical results for Pd in OREAS 235b (Indicative Value 0.719 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.34 | ppb
|  |  | 2.00: FA*MS | 3
|  |  | 2.00: 25g | 3
| 1 | 1 | 2.00: 0.70764 | 1
| 1 | 2 | 2.00: 0.7072 | 1
| 1 | 3 | 2.00: 0.73839 | 16
| 1 | 4 | 2.00: 0.70065 | 0.71898
| 1 | 5 | 2.00: 0.67859 | 7
| 1 | 6 | 2.00: 0.78141
| Mean |  | 2.00: 0.71898
| Median |  | 2.00: 0.70742
| Std Dev. |  | 2.00: 0.03608
| Rel.Std.Dev. |  | 2.00: 5.018%
| PDM3 |  | 2.00: 2.22e-14%
Z-Score (Absolute) | bar: 2.00
| Analytical results for Pt in OREAS 235b (Indicative Value 0.522 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.34 | ppb
|  |  | 2.00: FA*MS | 3
|  |  | 2.00: 25g | 3
| 1 | 1 | 2.00: 0.4748 | 1
| 1 | 2 | 2.00: 0.49878 | 1
| 1 | 3 | 2.00: 0.58139 | 16
| 1 | 4 | 2.00: 0.52767 | 0.52246
| 1 | 5 | 2.00: 0.52029 | 7
| 1 | 6 | 2.00: 0.53185
| Mean |  | 2.00: 0.52246
| Median |  | 2.00: 0.52398
| Std Dev. |  | 2.00: 0.035847
| Rel.Std.Dev. |  | 2.00: 6.861%
| PDM3 |  | 2.00: 8.882e-14%
Z-Score (Absolute) | bar: 2.00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 235b (Certified Value 1.6 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 0.085g | 25g | 50g | 30g | 40g | 30g | 30g | 25g | 30g | 25g | 25g | 25g | 25g | 15g | 50g | 25g | 50g | 25g | 10g | 25g | 20g | 15g | 15g | 50g | 3
| 1 | 1 | 1.8006 | 1.61 | 1.69 | 1.6699 | 1.47 | 1.69 | 1.67 | > 0.5 | 1.57 | 1.53 | 1.64 | 1.5 | 1.66 | 1.6524 | 1.55 | 1.71 | 1.52 | 1.24 | 1.62 | 1.63 | 1.6421 | 1.4832 | 1.4842 | 1.5363 | 1
| 1 | 2 | 1.822 | 1.6 | 1.64 | 1.6408 | 1.5 | 1.75 | 1.64 | > 0.5 | 1.57 | 1.54 | 1.6 | 1.51 | 1.64 | 1.6261 | 1.76 | 1.71 | 1.52 | 1.16 | 1.62 | 1.63 | 1.6471 | 1.5404 | 1.4687 | 1.54 | 3
| 1 | 3 | 1.7607 | 1.61 | 1.7 | 1.6796 | 1.5 | 1.65 | 1.63 | > 0.5 | 1.56 | 1.54 | 1.63 | 1.57 | 1.68 | 1.65 | 1.56 | 1.67 | 1.55 | 1.23 | 1.61 | 1.666 | 1.6507 | 1.5228 | 1.4682 | 1.53 | 16
| 1 | 4 | 1.7099 | 1.6 | 1.64 | 1.699 | 1.51 | 1.68 | 1.67 | > 0.5 | 1.56 | 1.52 | 1.68 | 1.63 | 1.65 | 1.6227 | 1.77 | 1.66 | 1.51 | 1.16 | 1.62 | 1.608 | 1.659 | 1.4786 | 1.4671 | 1.54 | 1.6042
| 1 | 5 | 1.7833 | 1.62 | 1.69 | 1.6893 | 1.46 | 1.69 | 1.65 | > 0.5 | 1.52 | 1.52 | 1.71 | 1.55 | 1.64 | 1.604 | 1.7 | 1.7 | 1.56 | 1.19 | 1.63 | 1.64 | 1.6641 | 1.5008 | 1.4361 | 1.53 | 9
| 1 | 6 | 1.7824 | 1.66 | 1.67 | 1.6602 | 1.55 | 1.8 | 1.68 | > 0.5 | 1.59 | 1.54 | 1.61 | 1.47 | 1.65 | 1.6348 | 1.7 | 1.68 | 1.51 | 1.2 | 1.61 | 1.667 | 1.6827 | 1.4522 | 1.4823 | 1.5606
|  | 7 | 1.7649
|  | 8 | 1.838
|  | 9 | 1.8844
|  | 10 | 1.8058
|  | 11 | 1.8997
|  | 12 | 1.8108
|  | 13 | 1.686
|  | 14 | 1.7992
|  | 15 | 1.7051
|  | 16 | 1.6971
|  | 17 | 1.7463
|  | 18 | 1.7545
|  | 19 | 1.8733
|  | 20 | 1.7944
| Mean |  | 1.7859 | 1.6167 | 1.6717 | 1.6731 | 1.4983 | 1.71 | 1.6567 |  | 1.5617 | 1.5317 | 1.645 | 1.5383 | 1.6533 | 1.6317 | 1.6733 | 1.6883 | 1.5283 | 1.1967 | 1.6183 | 1.6402 | 1.6576 | 1.4963 | 1.4678 | 1.5395
| Median |  | 1.7889 | 1.61 | 1.68 | 1.6747 | 1.5 | 1.69 | 1.66 |  | 1.565 | 1.535 | 1.635 | 1.53 | 1.65 | 1.6304 | 1.7 | 1.69 | 1.52 | 1.195 | 1.62 | 1.635 | 1.6548 | 1.492 | 1.4685 | 1.5381
| Std Dev. |  | 0.060516 | 0.022509 | 0.026394 | 0.020954 | 0.031885 | 0.054772 | 0.019664 |  | 0.023166 | 0.0098319 | 0.042308 | 0.057417 | 0.015055 | 0.01818 | 0.096264 | 0.02137 | 0.02137 | 0.033862 | 0.0075277 | 0.02293 | 0.014645 | 0.031928 | 0.017237 | 0.011282
| Rel.Std.Dev. |  | 3.388% | 1.392% | 1.579% | 1.252% | 2.128% | 3.203% | 1.187% |  | 1.483% | 0.6419% | 2.572% | 3.732% | 0.9106% | 1.114% | 5.753% | 1.266% | 1.398% | 2.83% | 0.4652% | 1.398% | 0.8835% | 2.134% | 1.174% | 0.7329%
| PDM3 |  | 11.32% | 0.7742% | 4.203% | 4.294% | -6.602% | 6.592% | 3.268% |  | -2.654% | -4.524% | 2.54% | -4.109% | 3.06% | 1.709% | 4.307% | 5.242% | -4.732% | -25.41% | 0.8781% | 2.239% | 3.327% | -6.727% | -8.507% | -4.037%
| Z-Score (Absolute) |  | NA | 0.1 | 0.56 | 0.58 | 1.53 | 1.03 | 0.38 | NA | 0.76 | 1.13 | 0.24 | 1.05 | 0.34 | 0.08 | 0.58 | 0.76 | 1.17 | 5.17 | 0.08 | 0.18 | 0.39 | 1.55 | 1.9 | 1.03
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 235b (Certified Value 1.43 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.28 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*OES | 2
|  |  | 0.085g | 200g | 200g | 200g | 50g | 50g | 20g | 200g | 05g | 30g | 30g | 30g | 30g | 30g | 200g | 50g | 30g | 15g | 3
| 1 | 1 | 1.8006 | 1.42 | 1.44 | 1.47 | 1.36 | 1.48 | 1.38 | 1.47 | 1.5 | 1.42 | 1.51 | 1.44 | 1.41 | 1.32 | 1.2955 | 1.62 | 1.37 | 1.4345 | 1
| 1 | 2 | 1.822 | 1.46 | 1.41 | 1.57 | 1.34 | 1.47 | 1.36 | 1.48 | 1.5 | 1.43 | 1.49 | 1.43 | 1.41 | 1.32 | 1.3232 | 1.55 | 1.36 | 1.425 | 3
| 1 | 3 | 1.7607 | 1.44 | 1.41 | 1.42 | 1.31 | 1.51 | 1.36 | 1.47 | 1.5 | 1.42 | 1.48 | 1.38 | 1.4 | 1.35 | 1.3085 | 1.593 | 1.38 | 1.4345 | 16
| 1 | 4 | 1.7099 | 1.42 | 1.46 | 1.51 | 1.37 | 1.48 | 1.35 | 1.47 | 1.5 | 1.42 | 1.47 | 1.45 | 1.42 | 1.32 | 1.2993 | 1.58 | 1.42 | 1.4345 | 1.4298
| 1 | 5 | 1.7833 | 1.43 | 1.56 | 1.53 | 1.34 | 1.49 | 1.41 | 1.46 | 1.5 | 1.45 | 1.47 | 1.48 | 1.4 | 1.34 | 1.3196 | 1.57 | 1.4 | 1.444 | 11
| 1 | 6 | 1.7824 | 1.43 | 1.49 | 1.59 | 1.33 | 1.49 | 1.4 | 1.46 | 1.5 | 1.41 | 1.47 | 1.46 | 1.41 | 1.27 | 1.3209 | 1.59 | 1.4 | 1.4345
|  | 7 | 2.00: 1.7649
|  | 8 | 2.00: 1.838
|  | 9 | 2.00: 1.8844
|  | 10 | 2.00: 1.8058
|  | 11 | 2.00: 1.8997
|  | 12 | 2.00: 1.8108
|  | 13 | 2.00: 1.686
|  | 14 | 2.00: 1.7992
|  | 15 | 2.00: 1.7051
|  | 16 | 2.00: 1.6971
|  | 17 | 2.00: 1.7463
|  | 18 | 2.00: 1.7545
|  | 19 | 2.00: 1.8733
|  | 20 | 2.00: 1.7944
| Mean |  | 1.7859 | 1.4333 | 1.4617 | 1.515 | 1.3417 | 1.4867 | 1.3767 | 1.4683 | 1.5 | 1.425 | 1.4817 | 1.44 | 1.4083 | 1.32 | 1.3112 | 1.5838 | 1.3883 | 1.4345
| Median |  | 1.7889 | 1.43 | 1.45 | 1.52 | 1.34 | 1.485 | 1.37 | 1.47 | 1.5 | 1.42 | 1.475 | 1.445 | 1.41 | 1.32 | 1.314 | 1.585 | 1.39 | 1.4345
| Std Dev. |  | 0.060516 | 0.015055 | 0.057067 | 0.063166 | 0.02137 | 0.013663 | 0.024221 | 0.0075277 | 0 | 0.013784 | 0.016021 | 0.034059 | 0.0075277 | 0.027568 | 0.011848 | 0.023583 | 0.022286 | 0.0060083
| Rel.Std.Dev. |  | 3.388% | 1.05% | 3.904% | 4.169% | 1.593% | 0.919% | 1.759% | 0.5127% | 0% | 0.9673% | 1.081% | 2.365% | 0.5345% | 2.088% | 0.9036% | 1.489% | 1.605% | 0.4188%
| PDM3 |  | 24.91% | 0.2499% | 2.232% | 5.962% | -6.161% | 3.98% | -3.713% | 2.698% | 4.913% | -0.333% | 3.63% | 0.7162% | -1.499% | -7.677% | -8.294% | 10.78% | -2.898% | 0.3315%
| Z-Score (Absolute) |  | NA | 0.01 | 0.4 | 1.17 | 1.33 | 0.76 | 0.83 | 0.5 | NA | 0.13 | 0.69 | 0.09 | 0.37 | 1.65 | 1.77 | 2.17 | 0.66 | 0.01
Results from laboratory 1.10 were removed due to their 0.1 ppm reading resolution.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 235b (Certified Value 1.68 ppm) | [6 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 1.8006 | 1.62 | 1.72 | 1.6 | 1.69 | 1
| 1 | 2 | 1.822 | 1.7 | 1.71 | 1.6 | 1.62 | 3
| 1 | 3 | 1.7607 | 1.51 | 1.73 | 1.6 | 1.63 | 16
| 1 | 4 | 1.7099 | 1.59 | 1.61 | 1.6 | 1.61 | 1.6772
| 1 | 5 | 1.7833 | 1.57 | 1.69 | 1.6 | 1.65 | 13
| 1 | 6 | 1.7824 | 1.65 | 1.62 | 1.6 | 1.76
| 2 | 7 | 1.7649 | 1.703 | 1.73 | 1.66 | 1.71
| 2 | 8 | 1.838 | 1.686 | 1.56 | 1.66 | 1.74
| 2 | 9 | 1.8844 | 1.62 | 1.74 | 1.66 | 1.71
| 2 | 10 | 1.8058 | 1.682 | 1.65 | 1.66 | 1.72
| 2 | 11 | 1.8997 | 1.744 | 1.65 | 1.66 | 1.74
| 2 | 12 | 1.8108 | 1.755 | 1.71 | 1.66 | 1.74
| 3 | 13 | 1.686 |  | 1.65 | 1.63
| 3 | 14 | 1.7992 |  | 1.65 | 1.63
| 3 | 15 | 1.7051 |  | 1.53 | 1.63
| 3 | 16 | 1.6971 |  | 1.66 | 1.63
| 3 | 17 | 1.7463 |  | 1.73 | 1.63
| 3 | 18 | 1.7545 |  | 1.76 | 1.63
| 4 | 19 | 1.8733 |  | 1.65 | 1.634
| 4 | 20 | 1.7944 |  | 1.662 | 1.788
| 4 | 21 |  |  | 1.696 | 1.682
| 4 | 22 |  |  | 1.646 | 1.731
| 4 | 23 |  |  | 1.703 | 1.711
| 4 | 24 |  |  | 1.791 | 1.674
| 5 | 25 |  |  | 1.794 | 1.616
| 5 | 26 |  |  | 1.66 | 1.683
| 5 | 27 |  |  | 1.798 | 1.686
| 5 | 28 |  |  | 1.706 | 1.723
| 5 | 29 |  |  | 1.697 | 1.78
| 5 | 30 |  |  | 1.615 | 1.773
| 6 | 31 |  |  | 1.7422 | 1.663
| 6 | 32 |  |  | 1.7911 | 1.701
| 6 | 33 |  |  | 1.6735 | 1.721
| 6 | 34 |  |  | 1.7032 | 1.753
| 6 | 35 |  |  | 1.6849 | 1.801
| 6 | 36 |  |  | 1.7866 | 1.713
| Mean |  | 1.7859 | 1.6525 | 1.6917 | 1.6715 | 1.6933
| Median |  | 1.7889 | 1.666 | 1.6965 | 1.66 | 1.71
| Std Dev. |  | 0.060516 | 0.073122 | 0.063626 | 0.058229 | 0.052628
| Rel.Std.Dev. |  | 3.388% | 4.425% | 3.761% | 3.484% | 3.108%
| PDM3 |  | 6.48% | -1.475% | 0.8594% | -0.3439% | 0.9595%
| Z-Score (Absolute) |  | NA | 1.79 | 0.62 | 0.62 | 0.73
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 235b (Certified Value 0.135 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.13 | < 0.2 | 0.19 | 0.66575 | 0.2 | < 0.5 | 0.3 | 0.3 | 0.14 | 0.15 | 0.16 | 0.14 | 0.09 | < 1 | 0.32 | 0.25 | < 0.5 | 0.07 | 0.16 | 0.147 | 0.32779 | 1.5171 |  |  |  |  | 1
| 1 | 2 | 0.16 | < 0.2 | 0.17 | 0.67243 | 0.2 | < 0.5 | 0.2 | 0.3 | 0.12 | 0.13 | 0.15 | 0.16 | < 0.02 | < 1 | 0.32 | 0.25 | < 0.5 | 0.08 | 0.16 | 0.136 | 0.28922 | 1.256 |  |  |  |  | 25
| 1 | 3 | 0.15 | < 0.2 | 0.16 | 0.6689 | 0.2 | < 0.5 | 0.3 | 0.3 | 0.13 | 0.13 | 0.15 | 0.17 | 0.1 | < 1 | 0.31 | 0.3 | < 0.5 | 0.08 | 0.15 | 0.122 | 0.31638 | 1.7123 |  |  |  |  | 16
| 1 | 4 | 0.14 | < 0.2 | 0.16 | 0.656 | 0.2 | < 0.5 | 0.2 | 0.3 | 0.13 | 0.19 | 0.16 | 0.15 | 0.08 | < 1 | 0.31 | 0.24 | < 0.5 | 0.13 | 0.18 | 0.124 | 0.25807 | 1.2199 |  |  |  |  | 0.13537
| 1 | 5 | 0.15 | < 0.2 | 0.15 | 0.6679 | 0.1 | < 0.5 | 0.3 | 0.2 | 0.14 | 0.13 | 0.16 | 0.16 | 0.06 | < 1 | 0.32 | 0.25 | < 0.5 | 0.09 | 0.15 | 0.141 | 0.27891 | 1.4619 |  |  |  |  | 15
| 1 | 6 | 0.15 | < 0.2 | 0.15 | 0.6655 | 0.2 | < 0.5 | 0.2 | 0.3 | 0.13 | 0.12 | 0.15 | 0.15 | 0.24 | < 1 | 0.33 | 0.24 | < 0.5 | 0.12 | 0.16 | 0.125 | 0.3685 | 1.135
| Mean |  | 0.14667 |  | 0.16333 | 0.66608 | 0.18333 |  | 0.25 | 0.28333 | 0.13167 | 0.14167 | 0.155 | 0.155 | 0.114 |  | 0.31833 | 0.255 |  | 0.095 | 0.16 | 0.1325 | 0.30648 | 1.3837
| Median |  | 0.15 |  | 0.16 | 0.66683 | 0.2 |  | 0.25 | 0.3 | 0.13 | 0.13 | 0.155 | 0.155 | 0.09 |  | 0.32 | 0.25 |  | 0.085 | 0.16 | 0.1305 | 0.3028 | 1.359
| Std Dev. |  | 0.010328 |  | 0.015055 | 0.005544 | 0.040825 |  | 0.054772 | 0.040825 | 0.0075277 | 0.025626 | 0.0054772 | 0.010488 | 0.071972 |  | 0.0075277 | 0.022583 |  | 0.02429 | 0.010954 | 0.01033 | 0.039497 | 0.21766
| Rel.Std.Dev. |  | 7.042% |  | 9.218% | 0.8323% | 22.27% |  | 21.91% | 14.41% | 5.717% | 18.09% | 3.534% | 6.767% | 63.13% |  | 2.365% | 8.856% |  | 25.57% | 6.847% | 7.796% | 12.89% | 15.73%
| PDM3 |  | 8.348% |  | 20.66% | 392.1% | 35.43% |  | 84.68% | 109.3% | -2.733% | 4.654% | 14.5% | 14.5% | -15.78% |  | 135.2% | 88.38% |  | -29.82% | 18.2% | -2.118% | 126.4% | 922.2%
| Z-Score (Absolute) |  | 0.14 | 0.64 | 0.04 | 5.39 | NA | 0.96 | NA | NA | 0.3 | 0.2 | 0.05 | 0.05 | 0.67 | 3.62 | 1.69 | 1.01 | 0.96 | 0.69 | 0 | 0.29 | 1.56 | 13.03
Results from laboratories 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 235b (Certified Value 7.57 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 7.6679 | 7.5508 | 7.76 | 7.17 | 7.37 | 7.914 | 5.25 | 7.84 | 7.22 | 7.53 | 7.08 | 7.57 | 7.46 | 7.3265 | 7.1865 | 7.621 | 7.59 | 7.68 | 7.33 | 7.63 | 6.92 | 7.3256 | 8.0929 | 7.5107 |  |  | 1
| 1 | 2 | 7.4589 | 7.5555 | 7.79 | 7 | 7.4005 | 7.895 | 6.24 | 7.59 | 7.14 | 7.47 | 7.78 | 7.65 | 7.52 | 7.2336 | 7.2999 | 7.621 | 7.7 | 7.62 | 7.28 | 7.75 | 7.35 | 7.2589 | 8.0396 | 7.479 |  |  | 0
| 1 | 3 | 7.5931 | 7.5483 | 7.8 | 6.97 | 7.3932 | 7.708 | 6.02 | 7.72 | 7.34 | 7.34 | 7.35 | 7.52 | 7.71 | 7.3962 | 7.4702 | 7.674 | 7.69 | 7.65 | 7.21 | 7.7 | 7.22 | 7.4185 | 7.9069 | 7.4702 |  |  | 16
| 1 | 4 | 7.5773 | 7.5395 | 7.7 | 6.86 | 7.41 | 7.929 | 7.9 | 7.66 | 7.2 | 7.51 | 7.78 | 7.59 | 7.65 | 7.37 | 7.4499 | 7.568 | 7.7 | 7.68 | 7.46 | 7.75 | 7.34 | 7.254 | 7.8449 | 7.4744 |  |  | 7.5667
| 1 | 5 | 7.594 | 7.5092 | 7.77 | 6.76 | 7.399 | 7.764 | 8.63 | 7.71 | 7.11 | 7.32 | 7.83 | 7.78 | 7.69 | 7.3791 | 7.3645 | 7.515 | 7.61 | 7.69 | 7.45 | 7.67 | 7.1 | 7.3263 | 7.7017 | 7.4832 |  |  | 16
| 1 | 6 | 7.7584 | 7.5244 | 7.78 | 6.55 | 7.372 | 7.754 | 8.3 | 7.57 | 7.24 | 7.26 | 7.7 | 7.77 | 7.65 | 7.6669 | 7.2835 | 7.621 | 7.64 | 7.66 | 7.37 | 7.79 | 7.19 | 7.3969 | 7.7651 | 7.4702
| Mean |  | 7.6083 | 7.5379 | 7.7667 | 6.885 | 7.3908 | 7.8273 | 7.0567 | 7.6817 | 7.2083 | 7.405 | 7.5867 | 7.6467 | 7.6133 | 7.3954 | 7.3424 | 7.6033 | 7.655 | 7.6633 | 7.35 | 7.715 | 7.1867 | 7.33 | 7.8918 | 7.4813
| Median |  | 7.5935 | 7.5439 | 7.775 | 6.915 | 7.3961 | 7.8295 | 7.07 | 7.685 | 7.21 | 7.405 | 7.74 | 7.62 | 7.65 | 7.3746 | 7.3322 | 7.621 | 7.665 | 7.67 | 7.35 | 7.725 | 7.205 | 7.3259 | 7.8759 | 7.4767
| Std Dev. |  | 0.099799 | 0.017839 | 0.03559 | 0.21455 | 0.016259 | 0.095974 | 1.3956 | 0.098675 | 0.081097 | 0.11256 | 0.30316 | 0.10783 | 0.10013 | 0.14529 | 0.10765 | 0.054738 | 0.048477 | 0.02582 | 0.097365 | 0.059245 | 0.1612 | 0.068077 | 0.15291 | 0.015279
| Rel.Std.Dev. |  | 1.312% | 0.2367% | 0.4582% | 3.116% | 0.22% | 1.226% | 19.78% | 1.285% | 1.125% | 1.52% | 3.996% | 1.41% | 1.315% | 1.965% | 1.466% | 0.7199% | 0.6333% | 0.3369% | 1.325% | 0.7679% | 2.243% | 0.9287% | 1.938% | 0.2042%
| PDM3 |  | 0.5493% | -0.38% | 2.643% | -9.009% | -2.325% | 3.444% | -6.74% | 1.519% | -4.736% | -2.137% | 0.2639% | 1.057% | 0.6163% | -2.264% | -2.964% | 0.4842% | 1.167% | 1.277% | -2.864% | 1.96% | -5.022% | -3.128% | 4.297% | -1.129%
| Z-Score (Absolute) |  | 0.19 | 0.1 | 0.85 | 2.82 | 0.71 | 1.1 | 2.1 | 0.5 | 1.47 | 0.65 | 0.1 | 0.35 | 0.21 | 0.69 | 0.91 | 0.17 | 0.39 | 0.42 | 0.88 | 0.64 | 1.56 | 0.97 | 1.37 | 0.34
| Analytical results for As in OREAS 235b (Certified Value 353 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 0
0
| 1 | 1 | 352.9 | 329 | 356 | 366.8558 | 349 | 358 | 388 | 50.9 | 353 | 377 | 359 | 352 | 361.8 | 329.8944 | 334 | 350.2 | 357.6 | 373.9 | 194 | 338.6 | 647.7631 | [5 empty] | 1
| 1 | 2 | 354 | 338 | 333 | 366.8599 | 351 | 346 | 357 | 48.3 | 350 | 358 | 358 | 358 | 363.9 | 329.5761 | 342 | 354.9 | 357.3 | 380.4 | 214.2 | 336.2 | 675.1901 | [5 empty] | 26
| 1 | 3 | 362.1 | 329 | 303 | 366 | 349 | 360 | 320 | 48.5 | 341 | 371 | 339 | 376 | 361 | 339.1221 | 375 | 355.8 | 344.7 | 374.6 | 197.1 | 343.6 | 742.8838 | [5 empty] | 16
| 1 | 4 | 362.4 | 330 | 305 | 367.667 | 353 | 363 | 346 | 48.2 | 354 | 380 | 341 | 370 | 369 | 325.0689 | 351 | 351.2 | 350.9 | 380 | 184.7 | 349.4 | 761.7373 | [5 empty] | 353.2216
| 1 | 5 | 358.4 | 341 | 291 | 368.0155 | 353 | 377 | 335 | 48.7 | 349 | 379 | 352 | 374 | 366.7 | 338.7791 | 352 | 358.2 | 369.1 | 377.1 | 199.4 | 342.6 | 697.1316 | [5 empty] | 17
| 1 | 6 | 364.9 | 347 | 284 | 365.31 | 351 | 364 | 353 | 50.7 | 343 | 358 | 343 | 371 | 373.9 | 333.2798 | 329 | 353.4 | 360.8 | 378.3 | 189.5 | 340.2 | 679.1213
| Mean |  | 359.1167 | 335.6667 | 312 | 366.7847 | 351 | 361.3333 | 349.8333 | 49.2167 | 348.3333 | 370.5 | 348.6667 | 366.8333 | 366.05 | 332.6201 | 347.1667 | 353.95 | 356.7333 | 377.3833 | 196.4833 | 341.7667 | 700.6379
| Median |  | 360.25 | 334 | 304 | 366.8579 | 351 | 361.5 | 349.5 | 48.6 | 349.5 | 374 | 347.5 | 370.5 | 365.3 | 331.5871 | 346.5 | 354.15 | 357.45 | 377.7 | 195.55 | 341.4 | 688.1264
| Std Dev. |  | 4.8676 | 7.5277 | 27.3203 | 1.0094 | 1.7889 | 10.0333 | 22.9906 | 1.24 | 5.2789 | 10.1735 | 8.8242 | 9.6003 | 4.8772 | 5.5563 | 16.388 | 2.9771 | 8.3591 | 2.711 | 10.1641 | 4.5999 | 43.4429
| Rel.Std.Dev. |  | 1.355% | 2.243% | 8.757% | 0.2752% | 0.5096% | 2.777% | 6.572% | 2.52% | 1.515% | 2.746% | 2.531% | 2.617% | 1.332% | 1.67% | 4.721% | 0.8411% | 2.343% | 0.7184% | 5.173% | 1.346% | 6.2%
| PDM3 |  | 1.669% | -4.97% | -11.67% | 3.84% | -0.6289% | 2.297% | -0.9592% | -86.07% | -1.384% | 4.892% | -1.29% | 3.854% | 3.632% | -5.832% | -1.714% | 0.2062% | 0.9942% | 6.84% | -44.37% | -3.243% | 98.36%
| Z-Score (Absolute) |  | 0.36 | 0.69 | 1.75 | 0.71 | 0 | 0.46 | 0.05 | 13.52 | 0.12 | 0.87 | 0.1 | 0.71 | 0.67 | 0.82 | 0.17 | 0.13 | 0.26 | 1.18 | 6.92 | 0.41 | 15.67
| Analytical results for Au in OREAS 235b (Indicative Value 1.71 ppm) | [27 empty] | Indicative
Au | Round | Replicate | 1.01: Lab | [5 empty] | 1
| No | No | 1.01: 1.30 | [5 empty] | ppm
|  |  | 1.01: 4A*MS | [5 empty] | 2
2
| 1 | 1 | 1.01: 2.2873 | [5 empty] | 1
| 1 | 2 | 1.01: 1.7653 | [5 empty] | 3
| 1 | 3 | 1.01: 1.454 | [5 empty] | 16
| 1 | 4 | 1.01: 1.3878 | [5 empty] | 1.7112
| 1 | 5 | 1.01: 1.7386 | [5 empty] | 9
| 1 | 6 | 1.01: 1.6341
| Mean |  | 1.01: 1.7112
| Median |  | 1.01: 1.6864
| Std Dev. |  | 1.01: 0.32008
| Rel.Std.Dev. |  | 1.01: 18.71%
| PDM3 |  | 1.01: 1.998e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for B in OREAS 235b (Indicative Value 30.6 ppm) | [27 empty] | Indicative
B | Round | Replicate | 1.01: Lab | [5 empty] | 1
| No | No | 1.01: 1.07 | [5 empty] | ppm
|  |  | 1.01: 4A*OES | [5 empty] | 1
1
| 1 | 1 | 1.01: 27.66 | [5 empty] | 1
| 1 | 2 | 1.01: 33.32 | [5 empty] | 4
| 1 | 3 | 1.01: 27.88 | [5 empty] | 16
| 1 | 4 | 1.01: 32.01 | [5 empty] | 30.5583
| 1 | 5 | 1.01: 30.66 | [5 empty] | 10
| 1 | 6 | 1.01: 31.82
| Mean |  | 1.01: 30.5583
| Median |  | 1.01: 31.24
| Std Dev. |  | 1.01: 2.3199
| Rel.Std.Dev. |  | 1.01: 7.592%
| PDM3 |  | 1.01: 1.11e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ba in OREAS 235b (Certified Value 691 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 693.1 | 693.2222 | 704 | 663 | 657.17 | 724 | 695 | 787 | 672 | 700 | 650 | 680 | 690 | 679 | 682.429 | 691 | 607.9 | 684 | 689.3 | 704 | 676 | 700 | 655.4 | 719 | 26.926 | 610.4791 | 1
| 1 | 2 | 679.5 | 689.0327 | 720 | 703 | 656.448 | 725 | 693 | 722 | 661 | 700 | 710 | 690 | 680 | 680 | 674.5858 | 696 | 648.6 | 700 | 690.1 | 712 | 687 | 664 | 650.7 | 711 | 727.168 | 598.6555 | 27
| 1 | 3 | 683.4 | 691.1087 | 688 | 727 | 653.2 | 721 | 699 | 696 | 679 | 680 | 680 | 680 | 700 | 679 | 681.385 | 691 | 612 | 694 | 692.1 | 697 | 681 | 685 | 664.1 | 700 | 709.9848 | 611.8318 | 16
| 1 | 4 | 690.8 | 688.7143 | 697 | 725 | 655.26 | 741 | 724 | 691 | 661 | 700 | 720 | 690 | 700 | 684 | 687.1485 | 692 | 624.5 | 701 | 691.4 | 691 | 687 | 646 | 649.6 | 698 | 766.0022 | 593.7683 | 691.2618
| 1 | 5 | 687.1 | 692.4604 | 701 | 743 | 654.26 | 720 | 735 | 703 | 655 | 670 | 720 | 700 | 700 | 686 | 680.3622 | 688 | 629.8 | 682 | 693.5 | 689 | 681 | 656 | 655.4 | 688 | 700.8531 | 609.167 | 18
| 1 | 6 | 693.7 | 693.8618 | 724 | 712 | 654.3068 | 741 | 734 | 732 | 653 | 680 | 700 | 700 | 690 | 699 | 677.5506 | 698 | 660.7 | 671 | 698.4 | 683 | 693 | 655 | 659.8 | 691 | 708.6592 | 588.7249
| Mean |  | 687.9333 | 691.4 | 705.6667 | 712.1667 | 655.1075 | 728.6667 | 713.3333 | 721.8333 | 663.5 | 688.3333 | 696.6667 | 690 | 693.3333 | 684.5 | 680.5769 | 692.6667 | 630.5833 | 688.6667 | 692.4667 | 696 | 684.1667 | 667.6667 | 655.8333 | 701.1667 | 606.5989 | 602.1044
| Median |  | 688.95 | 691.7846 | 702.5 | 718.5 | 654.7834 | 724.5 | 711.5 | 712.5 | 661 | 690 | 705 | 690 | 695 | 682 | 680.8736 | 691.5 | 627.15 | 689 | 691.75 | 694 | 684 | 660 | 655.4 | 699 | 709.322 | 603.9112
| Std Dev. |  | 5.6663 | 2.1637 | 13.8082 | 27.7014 | 1.4881 | 9.7297 | 19.8259 | 35.5607 | 10.0747 | 13.2916 | 27.3252 | 8.9443 | 8.165 | 7.6616 | 4.2981 | 3.6697 | 20.6354 | 11.7246 | 3.2611 | 10.6207 | 6.0139 | 20.6171 | 5.4709 | 11.8561 | 284.9406 | 9.7472
| Rel.Std.Dev. |  | 0.8237% | 0.3129% | 1.957% | 3.89% | 0.2272% | 1.335% | 2.779% | 4.926% | 1.518% | 1.931% | 3.922% | 1.296% | 1.178% | 1.119% | 0.6315% | 0.5298% | 3.272% | 1.703% | 0.4709% | 1.526% | 0.879% | 3.088% | 0.8342% | 1.691% | 46.97% | 1.619%
| PDM3 |  | -0.4815% | 0.01999% | 2.084% | 3.024% | -5.23% | 5.411% | 3.193% | 4.423% | -4.016% | -0.4236% | 0.7819% | -0.1825% | 0.2997% | -0.9782% | -1.546% | 0.2032% | -8.778% | -0.3754% | 0.1743% | 0.6854% | -1.026% | -3.413% | -5.125% | 1.433% | -12.25% | -12.9%
| Z-Score (Absolute) |  | 0.09 | 0.14 | 1.07 | 1.5 | 2.24 | 2.58 | 1.57 | 2.13 | 1.69 | 0.07 | 0.48 | 0.04 | 0.26 | 0.32 | 0.57 | 0.22 | 3.85 | 0.04 | 0.21 | 0.44 | 0.34 | 1.42 | 2.19 | 0.78 | 5.42 | 5.71
| Analytical results for Be in OREAS 235b (Certified Value 2.54 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.46 | 3 | 2.4 | < 5 | 2.3 | 2.55 | 2.8 | 2 | 2.43 | 2.4 | 2.66 | 2.31 | 3 | 2.5426 | 3.1 | 2.135 | 2.7 | 3.44 | 2.24 | 2.46 | 3 | 2.4464 | 2.8867 | 2.5418 |  |  | 1
| 1 | 2 | 2.45 | 3 | 2.4 | < 5 | 2.2 | 2.47 | 2.8 | 2 | 2.35 | 2.62 | 2.48 | 2.33 | 3 | 2.5188 | 3.2 | 2.291 | 2.7 | 3.38 | 2.63 | 2.43 | 3 | 2.431 | 2.9424 | 2.5446 |  |  | 28
| 1 | 3 | 2.63 | 3 | 2.6 | < 5 | 2.3 | 2.46 | 2.5 | 3 | 2.35 | 2.48 | 2.48 | 2.4 | 3 | 2.4683 | 3.1 | 2.118 | 3 | 3.43 | 2.9 | 2.58 | 3 | 2.4783 | 2.8867 | 2.5598 |  |  | 16
| 1 | 4 | 2.63 | 3 | 2.4 | < 5 | 2.3 | 2.6 | 2.4 | 2 | 2.46 | 2.63 | 2.6 | 2.41 | 3 | 2.4961 | 3 | 2.198 | 2.9 | 3.39 | 2.52 | 2.47 | 3 | 2.4288 | 2.7199 | 2.4772 |  |  | 2.5425
| 1 | 5 | 2.5 | 3 | 2.4 | < 5 | 2.2 | 2.64 | 2.4 | 2 | 2.41 | 2.82 | 2.71 | 2.4 | 3 | 2.5026 | 3 | 2.194 | 2.8 | 3.49 | 2.49 | 2.47 | 3 | 2.4519 | 2.998 | 2.4769 |  |  | 19
| 1 | 6 | 2.65 | 3 | 2.2 | < 5 | 2.4 | 2.65 | 2.8 | 2 | 2.44 | 2.67 | 2.5 | 2.47 | 3 | 2.5517 | 3 | 2.251 | 2.6 | 3.36 | 2.55 | 2.35 | 3 | 2.4695 | 2.9424 | 2.5472
| Mean |  | 2.5533 | 3 | 2.4 |  | 2.2833 | 2.5617 | 2.6167 | 2.1667 | 2.4067 | 2.6033 | 2.5717 | 2.3867 | 3 | 2.5134 | 3.0667 | 2.1978 | 2.7833 | 3.415 | 2.555 | 2.46 | 3 | 2.451 | 2.896 | 2.5246
| Median |  | 2.565 | 3 | 2.4 |  | 2.3 | 2.575 | 2.65 | 2 | 2.42 | 2.625 | 2.55 | 2.4 | 3 | 2.5107 | 3.05 | 2.196 | 2.75 | 3.41 | 2.535 | 2.465 | 3 | 2.4491 | 2.9146 | 2.5432
| Std Dev. |  | 0.093095 | 0 | 0.12649 |  | 0.075277 | 0.082805 | 0.20412 | 0.40825 | 0.046762 | 0.1476 | 0.099683 | 0.058195 | 0 | 0.030993 | 0.08165 | 0.066095 | 0.1472 | 0.047645 | 0.21417 | 0.074297 | 0 | 0.020012 | 0.095806 | 0.037325
| Rel.Std.Dev. |  | 3.646% | 0% | 5.27% |  | 3.297% | 3.232% | 7.801% | 18.84% | 1.943% | 5.67% | 3.876% | 2.438% | 0% | 1.233% | 2.662% | 3.007% | 5.288% | 1.395% | 8.382% | 3.02% | 0% | 0.8165% | 3.308% | 1.478%
| PDM3 |  | 0.4273% | 18% | -5.604% |  | -10.19% | 0.7551% | 2.918% | -14.78% | -5.341% | 2.394% | 1.148% | -6.128% | 18% | -1.145% | 20.62% | -13.56% | 9.474% | 34.32% | 0.4929% | -3.244% | 18% | -3.598% | 13.91% | -0.7042%
| Z-Score (Absolute) |  | 0 | 2.94 | 1.01 | 0.35 | 1.78 | 0.05 | 0.42 | NA | 0.97 | 0.33 | 0.12 | 1.1 | NA | 0.26 | 3.38 | 2.34 | 1.52 | 5.68 | 0.01 | 0.61 | NA | 0.67 | 2.26 | 0.19
Results from laboratories 1.11, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 235b (Certified Value 0.34 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 0.41 | 0.37764 | 0.3 | 0.35 | < 5 | 0.35 | 0.29 | 0.38 | 0.4 | 0.35 | 0.34 | 0.35 | 0.31 | 0.35 | 0.34806 | 0.4 | 0.329 | 0.36 | < 5 | 0.35 | 0.31 | 0.4 | 0.5431 | < 5 | 1.1163 | 0.35542 | 1
| 1 | 2 | 0.42 | 0.35661 | 0.3 | 0.33 | < 5 | 0.32 | 0.28 | 0.38 | 0.5 | 0.33 | 0.34 | 0.36 | 0.31 | 0.35 | 0.34455 | 0.6 | 0.331 | 0.35 | < 5 | 0.44 | 0.34 | 0.36 | 0.5011 | < 5 | 1.2286 | 0.36407 | 29
| 1 | 3 | 0.35 | 0.36869 | 0.3 | 0.35 | < 5 | 0.34 | 0.32 | 0.29 | 0.4 | 0.34 | 0.44 | 0.35 | 0.33 | 0.34 | 0.34732 | 0.5 | 0.305 | 0.32 | < 5 | 0.38 | 0.3 | 0.36 | 0.4242 | < 5 | 1.069 | 0.33837 | 16
| 1 | 4 | 0.36 | 0.38358 | 0.4 | 0.33 | < 5 | 0.36 | 0.4 | 0.38 | 0.4 | 0.32 | 0.38 | 0.35 | 0.35 | 0.34 | 0.33031 | 0.4 | 0.34 | 0.35 | < 5 | 0.35 | 0.29 | 0.33 | 0.4551 | < 5 | 1.8412 | 0.35483 | 0.34364
| 1 | 5 | 0.33 | 0.3742 | 0.4 | 0.32 | < 5 | 0.34 | 0.37 | 0.46 | 0.4 | 0.32 | 0.34 | 0.36 | 0.31 | 0.35 | 0.35648 | 0.4 | 0.311 | 0.33 | < 5 | 0.4 | 0.28 | 0.34 | 0.443 | < 5 | 1.5359 | 0.34572 | 20
| 1 | 6 | 0.34 | 0.38023 | 0.4 | 0.29 | < 5 | 0.34 | 0.36 | 0.35 | 0.4 | 0.33 | 0.34 | 0.33 | 0.33 | 0.34 | 0.33792 | 0.5 | 0.335 | 0.31 | < 5 | 0.34 | 0.33 | 0.33 | 0.4473 | < 5 | 1.264 | 0.35005
| Mean |  | 0.36833 | 0.37349 | 0.35 | 0.32833 |  | 0.34167 | 0.33667 | 0.37333 | 0.41667 | 0.33167 | 0.36333 | 0.35 | 0.32333 | 0.345 | 0.34411 | 0.46667 | 0.32517 | 0.33667 |  | 0.37667 | 0.30833 | 0.35333 | 0.46897 |  | 1.3425 | 0.35141
| Median |  | 0.355 | 0.37592 | 0.35 | 0.33 |  | 0.34 | 0.34 | 0.38 | 0.4 | 0.33 | 0.34 | 0.35 | 0.32 | 0.345 | 0.34594 | 0.45 | 0.33 | 0.34 |  | 0.365 | 0.305 | 0.35 | 0.4512 |  | 1.2463 | 0.35244
| Std Dev. |  | 0.037639 | 0.0097264 | 0.054772 | 0.022286 |  | 0.013292 | 0.04761 | 0.055015 | 0.040825 | 0.01169 | 0.040825 | 0.010954 | 0.01633 | 0.0054772 | 0.0090323 | 0.08165 | 0.013949 | 0.019664 |  | 0.038297 | 0.023166 | 0.026583 | 0.044426 |  | 0.29367 | 0.008858
| Rel.Std.Dev. |  | 10.22% | 2.604% | 15.65% | 6.788% |  | 3.89% | 14.14% | 14.74% | 9.798% | 3.525% | 11.24% | 3.13% | 5.05% | 1.588% | 2.625% | 17.5% | 4.29% | 5.841% |  | 10.17% | 7.513% | 7.524% | 9.473% |  | 21.87% | 2.521%
| PDM3 |  | 7.186% | 8.687% | 1.851% | -4.454% |  | -0.574% | -2.029% | 8.641% | 21.25% | -3.484% | 5.731% | 1.851% | -5.909% | 0.396% | 0.136% | 35.8% | -5.376% | -2.029% |  | 9.611% | -10.27% | 2.821% | 36.47% |  | 290.7% | 2.261%
| Z-Score (Absolute) |  | 0.52 | 0.68 | NA | 0.71 | 66.09 | 0.3 | 0.45 | 0.67 | NA | 0.61 | 0.37 | 0.04 | 0.86 | 0.2 | 0.22 | NA | 0.81 | 0.45 | 66.09 | 0.78 | 1.32 | 0.06 | 3.62 | 66.09 | 30.48 | 0
Results from laboratories 1.05, 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 235b (Certified Value 0.86 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.8644 | 0.85228 | 0.89 | 0.88 | 0.83333 | 0.926 | 0.69 | 0.854 | 0.89 | 0.86 | 0.82 | 0.87 | 0.87 | 0.8027 | 0.83909 | 0.786 | 0.85 | 0.91 | 0.85 | 0.85 | 0.77 | 0.82974 | 0.9305 |  |  |  | 1
| 1 | 2 | 0.8412 | 0.8561 | 0.89 | 0.79 | 0.81 | 0.913 | 0.8 | 0.827 | 0.87 | 0.86 | 0.89 | 0.88 | 0.88 | 0.8078 | 0.84908 | 0.786 | 0.86 | 0.92 | 0.86 | 0.86 | 0.83 | 0.82476 | 0.9721 |  |  |  | 0
| 1 | 3 | 0.8516 | 0.85318 | 0.85 | 0.83 | 0.82333 | 0.912 | 0.67 | 0.851 | 0.86 | 0.83 | 0.84 | 0.87 | 0.9 | 0.8262 | 0.84494 | 0.786 | 0.86 | 0.91 | 0.85 | 0.85 | 0.82 | 0.8398 | 0.9047 |  |  |  | 16
| 1 | 4 | 0.866 | 0.85266 | 0.87 | 0.81 | 0.825 | 0.925 | 0.95 | 0.846 | 0.88 | 0.85 | 0.92 | 0.88 | 0.89 | 0.8384 | 0.84049 | 0.779 | 0.87 | 0.92 | 0.87 | 0.86 | 0.82 | 0.82053 | 0.9026 |  |  |  | 0.85969
| 1 | 5 | 0.8617 | 0.85749 | 0.88 | 0.78 | 0.84 | 0.904 | 1.02 | 0.847 | 0.86 | 0.84 | 0.9 | 0.9 | 0.9 | 0.8364 | 0.83747 | 0.779 | 0.85 | 0.91 | 0.86 | 0.84 | 0.81 | 0.83152 | 0.8917 |  |  |  | 21
| 1 | 6 | 0.8814 | 0.85858 | 0.88 | 0.8 | 0.85 | 0.908 | 0.98 | 0.825 | 0.89 | 0.85 | 0.89 | 0.9 | 0.9 | 0.865 | 0.84845 | 0.786 | 0.86 | 0.9 | 0.85 | 0.87 | 0.83 | 0.83425 | 0.8951
| Mean |  | 0.86105 | 0.85505 | 0.87667 | 0.815 | 0.83028 | 0.91467 | 0.85167 | 0.84167 | 0.875 | 0.84833 | 0.87667 | 0.88333 | 0.89 | 0.82942 | 0.84325 | 0.78367 | 0.85833 | 0.91167 | 0.85667 | 0.855 | 0.81333 | 0.8301 | 0.91612
| Median |  | 0.86305 | 0.85464 | 0.88 | 0.805 | 0.82917 | 0.9125 | 0.875 | 0.8465 | 0.875 | 0.85 | 0.89 | 0.88 | 0.895 | 0.8313 | 0.84272 | 0.786 | 0.86 | 0.91 | 0.855 | 0.855 | 0.82 | 0.83063 | 0.90365
| Std Dev. |  | 0.013665 | 0.0026981 | 0.015055 | 0.036194 | 0.014001 | 0.0089815 | 0.1525 | 0.012485 | 0.013784 | 0.01169 | 0.038297 | 0.013663 | 0.012649 | 0.022747 | 0.004943 | 0.0036148 | 0.0075277 | 0.0075277 | 0.008165 | 0.010488 | 0.022509 | 0.0068354 | 0.030633
| Rel.Std.Dev. |  | 1.587% | 0.3156% | 1.717% | 4.441% | 1.686% | 0.9819% | 17.91% | 1.483% | 1.575% | 1.378% | 4.368% | 1.547% | 1.421% | 2.742% | 0.5862% | 0.4613% | 0.877% | 0.8257% | 0.9531% | 1.227% | 2.768% | 0.8234% | 3.344%
| PDM3 |  | 0.1583% | -0.5398% | 1.975% | -5.198% | -3.421% | 6.395% | -0.9332% | -2.096% | 1.781% | -1.321% | 1.975% | 2.75% | 3.526% | -3.521% | -1.912% | -8.843% | -0.1577% | 6.046% | -0.3516% | -0.5454% | -5.392% | -3.442% | 6.564%
| Z-Score (Absolute) |  | 0.19 | 0 | 0.67 | 1.25 | 0.77 | 1.86 | 0.11 | 0.42 | 0.62 | 0.21 | 0.67 | 0.88 | 1.09 | 0.8 | 0.37 | 2.22 | 0.1 | 1.77 | 0.05 | 0 | 1.3 | 0.78 | 1.9
| Analytical results for Cd in OREAS 235b (Indicative Value 0.088 ppm) | [27 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.03 | 1.26 | < 0.5 | < 0.3 | 0.84 | 0.07 | 0.11 | 0.38 | 0.3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.1 | 0.10883 | < 0.1 | < 0.1 | 0.04 | 4.04 | 0.03 | 0.05 | 0.07 | 0.1625 | < 5 | 0.2331 | 0.05995 | 1
| 1 | 2 | 0.05 | 1.18 | < 0.5 | < 0.3 | 0.86031 | 0.06 | 0.1 | 0.4 | 0.2 | 0.04 | 0.04 | 0.05 | 0.04 | 0.12 | 0.12133 | < 0.1 | 0.109 | 0.06 | 4.49 | 0.03 | 0.04 | 0.05 | 0.1379 | < 5 | 0.2438 | 0.06131 | 5
| 1 | 3 | 0.06 | 1.31 | < 0.5 | 0.4 | 0.841 | 0.07 | 0.11 | 0.31 | 0.3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.11 | 0.12719 | 0.1 | < 0.1 | 0.05 | 4.05 | 0.06 | 0.05 | 0.06 | 0.1607 | < 5 | 0.2385 | 0.06698 | 16
| 1 | 4 | 0.06 | 1.2 | < 0.5 | < 0.3 | 0.8567 | 0.06 | 0.14 | 0.41 | 0.3 | 0.04 | 0.06 | 0.05 | 0.04 | 0.12 | 0.11031 | < 0.1 | 0.11 | 0.06 | 4.49 | 0.04 | 0.04 | 0.04 | 0.1807 | < 5 | 0.2456 | 0.06373 | 0.088371
| 1 | 5 | 0.07 | 1.32 | < 0.5 | 0.4 | 0.866 | 0.07 | 0.13 | 0.22 | 0.2 | 0.04 | 0.04 | 0.03 | 0.04 | 0.09 | 0.12093 | < 0.1 | < 0.1 | 0.04 | 4.61 | 0.07 | 0.05 | 0.04 | 0.1392 | < 5 | 0.2367 | 0.05835 | 11
| 1 | 6 | 0.05 | 1.39 | < 0.5 | < 0.3 | 0.86469 | 0.06 | 0.13 | 0.45 | 0.3 | 0.04 | 0.05 | 0.04 | 0.03 | 0.1 | 0.10794 | < 0.1 | < 0.1 | 0.05 | 4.15 | 0.03 | 0.05 | 0.05 | 0.1218 | < 5 | 0.2403 | 0.05966
| Mean |  | 0.053333 | 1.2767 |  | 0.4 | 0.85478 | 0.065 | 0.12 | 0.36167 | 0.26667 | 0.04 | 0.045 | 0.041667 | 0.038333 | 0.10667 | 0.11609 | 0.1 | 0.1095 | 0.05 | 4.305 | 0.043333 | 0.046667 | 0.051667 | 0.15047 |  | 0.23967 | 0.061663
| Median |  | 0.055 | 1.285 |  | 0.4 | 0.8585 | 0.065 | 0.12 | 0.39 | 0.3 | 0.04 | 0.04 | 0.04 | 0.04 | 0.105 | 0.11562 | 0.1 | 0.1095 | 0.05 | 4.32 | 0.035 | 0.05 | 0.05 | 0.14995 |  | 0.2394 | 0.06063
| Std Dev. |  | 0.013663 | 0.079162 |  | 0 | 0.011548 | 0.0054772 | 0.015492 | 0.083287 | 0.05164 | 0 | 0.0083666 | 0.0075277 | 0.0040825 | 0.012111 | 0.0080809 |  | 0.00070711 | 0.0089443 | 0.25328 | 0.017512 | 0.005164 | 0.01169 | 0.021308 |  | 0.0046038 | 0.0031809
| Rel.Std.Dev. |  | 25.62% | 6.201% |  | 0% | 1.351% | 8.427% | 12.91% | 23.03% | 19.36% | 0% | 18.59% | 18.07% | 10.65% | 11.35% | 6.961% |  | 0.6458% | 17.89% | 5.883% | 40.41% | 11.07% | 22.63% | 14.16% |  | 1.921% | 5.158%
| PDM3 |  | -39.65% | 1345% |  | 352.6% | 867.3% | -26.45% | 35.79% | 309.3% | 201.8% | -54.74% | -49.08% | -52.85% | -56.62% | 20.7% | 31.36% | 13.16% | 23.91% | -43.42% | 4771% | -50.96% | -47.19% | -41.53% | 70.27% |  | 171.2% | -30.22%
| Z-Score (Absolute) |  | 0.51 | 17.52 | 2.38 | 2.14 | 11.3 | 0.34 | 0.47 | 4.03 | 2.63 | 0.71 | 0.64 | 0.69 | 0.74 | 0.27 | 0.41 | 0.44 | 0.27 | 0.56 | 62.15 | 0.66 | 0.61 | 0.54 | 0.92 | 35.55 | 2.23 | 0.39
| Analytical results for Ce in OREAS 235b (Certified Value 83 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 0
1
| 1 | 1 | 85.08 | 91.4798 | 83.7 | 86.6 | 80.16 | 84 | 84.8 | 79.2 | 84.9 | 75.7 | 80.83 | 80.2196 | 69 | 77.45 | 76.8 | 86.78 | 85.04 | 84.05 | 71.8 | 67.0878 | 80.9297 | [5 empty] | 1
| 1 | 2 | 83.55 | 90.6252 | 87.2 | 87.9 | 83.8 | 83 | 84.4 | 85.2 | 84.1 | 76.9 | 83.05 | 77.2464 | 69 | 80.46 | 80.72 | 89.2 | 85.28 | 82.28 | 69.9 | 82.1326 | 80.8876 | [5 empty] | 31
| 1 | 3 | 85.76 | 93.4835 | 81.7 | 88.9 | 80.46 | 83 | 82.5 | 79.3 | 80.1 | 83.1 | 83.15 | 80.7331 | 68 | 80.51 | 81.31 | 89.45 | 81.81 | 84.13 | 70.1 | 82.5817 | 81.3501 | [5 empty] | 16
| 1 | 4 | 85.43 | 92.8127 | 82.9 | 87.7 | 94.72 | 83 | 85.4 | 86.4 | 83.4 | 81.2 | 83.81 | 78.2809 | 69 | 79.92 | 79.15 | 87.55 | 83.35 | 81.06 | 70.6 | 79.4829 | 79.637 | [5 empty] | 83.0499
| 1 | 5 | 86.53 | 93.0336 | 81.8 | 89.2 | 97.92 | 82 | 84.1 | 89.4 | 86.9 | 83.7 | 81.16 | 78.5805 | 66 | 81.21 | 76.38 | 86.57 | 81.22 | 80.6 | 70.7 | 80.426 | 81.5181 | [5 empty] | 22
| 1 | 6 | 87.73 | 91.7696 | 84.6 | 86.1 | 96.21 | 83 | 80.7 | 88.5 | 89.5 | 82.4 | 84.24 | 77.7597 | 69 | 80.76 | 77.13 | 84.57 | 85.68 | 81.26 | 71.5 | 83.2554 | 78.4124
| Mean |  | 85.68 | 92.2007 | 83.65 | 87.7333 | 88.8783 | 83 | 83.65 | 84.6667 | 84.8167 | 80.5 | 82.7067 | 78.8034 | 68.3333 | 80.0517 | 78.5817 | 87.3533 | 83.73 | 82.23 | 70.7667 | 79.1611 | 80.4558
| Median |  | 85.595 | 92.2911 | 83.3 | 87.8 | 89.26 | 83 | 84.25 | 85.8 | 84.5 | 81.8 | 83.1 | 78.4307 | 69 | 80.485 | 78.14 | 87.165 | 84.195 | 81.77 | 70.65 | 81.2793 | 80.9087
| Std Dev. |  | 1.4065 | 1.0864 | 2.0637 | 1.2242 | 8.274 | 0.63246 | 1.7421 | 4.4514 | 3.1953 | 3.3793 | 1.3998 | 1.3831 | 1.2111 | 1.3421 | 2.1198 | 1.8185 | 1.9001 | 1.5422 | 0.75277 | 6.0788 | 1.1988
| Rel.Std.Dev. |  | 1.642% | 1.178% | 2.467% | 1.395% | 9.309% | 0.762% | 2.083% | 5.258% | 3.767% | 4.198% | 1.692% | 1.755% | 1.772% | 1.677% | 2.698% | 2.082% | 2.269% | 1.875% | 1.064% | 7.679% | 1.49%
| PDM3 |  | 3.167% | 11.02% | 0.7225% | 5.639% | 7.018% | -0.06013% | 0.7225% | 1.947% | 2.127% | -3.07% | -0.4133% | -5.113% | -17.72% | -3.61% | -5.38% | 5.182% | 0.8189% | -0.9873% | -14.79% | -4.683% | -3.124%
| Z-Score (Absolute) |  | 0.67 | 2.31 | 0.16 | 1.19 | 1.48 | 0 | 0.16 | 0.42 | 0.46 | 0.63 | 0.07 | 1.06 | 3.69 | 0.74 | 1.11 | 1.1 | 0.18 | 0.19 | 3.08 | 0.97 | 0.64
| Analytical results for Co in OREAS 235b (Certified Value 17.8 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 17.7 | 15.2708 | 17 | 18 | 22.41 | 17.7 | 16 | 19 | 17.4 | 18.7 | 15.9 | 17.7 | 17.6 | 18.5 | 18.3273 | 18 | 15.11 | 16.8 | 24.72 | 18.1 | 18.3 | 18 | 20.425 | 22 | 20.3787 | 17.7792 | 1
| 1 | 2 | 17.6 | 15.6063 | 17 | 17 | 22.3533 | 17.7 | 17 | 18 | 16.8 | 18.2 | 17.3 | 17 | 17.6 | 18.6 | 17.6729 | 18 | 16.03 | 17.9 | 25.1 | 18.1 | 18.1 | 18.8 | 20.14 | 21 | 19.41 | 17.5342 | 32
| 1 | 3 | 17.6 | 15.1513 | 18 | 18 | 22.3567 | 17.6 | 17 | 17 | 17.9 | 17.7 | 16.8 | 16.7 | 17.9 | 18.5 | 17.4652 | 17.7 | 15.27 | 17.4 | 24.88 | 17.4 | 17.3 | 18.3 | 20.71 | 19 | 20.3453 | 17.5236 | 16
| 1 | 4 | 17.8 | 15.2963 | 17 | 18 | 22.99 | 17.9 | 16 | 17 | 16.7 | 18.6 | 18.8 | 17.4 | 18.2 | 18.1 | 17.6566 | 18.3 | 16.05 | 17.8 | 24.35 | 17.5 | 17.6 | 17.9 | 20.14 | 20 | 20.479 | 17.3555 | 17.7651
| 1 | 5 | 17.8 | 15.0803 | 16 | 18 | 22.6167 | 17.8 | 16 | 18 | 16.8 | 17.7 | 18.1 | 17.8 | 18.4 | 18.6 | 17.0232 | 17.9 | 15.4 | 18.1 | 24.2 | 17.2 | 17.2 | 17.7 | 19.76 | 20 | 21.1471 | 17.8001 | 23
| 1 | 6 | 17.7 | 15.619 | 17 | 18 | 22.3033 | 17.9 | 16 | 18 | 17 | 18.8 | 17.9 | 17.9 | 18 | 18.8 | 17.8061 | 17.7 | 15.78 | 16.6 | 24.74 | 17.4 | 17.9 | 18 | 20.045 | 21 | 20.479 | 17.6515
| Mean |  | 17.7 | 15.3373 | 17 | 17.8333 | 22.505 | 17.7667 | 16.3333 | 17.8333 | 17.1 | 18.2833 | 17.4667 | 17.4167 | 17.95 | 18.5167 | 17.6586 | 17.9333 | 15.6067 | 17.4333 | 24.665 | 17.6167 | 17.7333 | 18.1167 | 20.2033 | 20.5 | 20.3732 | 17.6074
| Median |  | 17.7 | 15.2836 | 17 | 18 | 22.3833 | 17.75 | 16 | 18 | 16.9 | 18.4 | 17.6 | 17.55 | 17.95 | 18.55 | 17.6648 | 17.95 | 15.59 | 17.6 | 24.73 | 17.45 | 17.75 | 18 | 20.14 | 20.5 | 20.4289 | 17.5929
| Std Dev. |  | 0.089443 | 0.22737 | 0.63246 | 0.40825 | 0.26168 | 0.12111 | 0.5164 | 0.75277 | 0.46476 | 0.49565 | 1.0289 | 0.47924 | 0.32094 | 0.23166 | 0.42667 | 0.22509 | 0.40213 | 0.61536 | 0.33453 | 0.38687 | 0.44121 | 0.38687 | 0.32726 | 1.0488 | 0.55675 | 0.16993
| Rel.Std.Dev. |  | 0.5053% | 1.482% | 3.72% | 2.289% | 1.163% | 0.6816% | 3.162% | 4.221% | 2.718% | 2.711% | 5.891% | 2.752% | 1.788% | 1.251% | 2.416% | 1.255% | 2.577% | 3.53% | 1.356% | 2.196% | 2.488% | 2.135% | 1.62% | 5.116% | 2.733% | 0.9651%
| PDM3 |  | -0.3662% | -13.67% | -4.307% | 0.3843% | 26.68% | 0.009042% | -8.059% | 0.3843% | -3.744% | 2.917% | -1.68% | -1.961% | 1.041% | 4.231% | -0.5995% | 0.9472% | -12.15% | -1.867% | 38.84% | -0.8353% | -0.1786% | 1.979% | 13.73% | 15.4% | 14.68% | -0.8877%
| Z-Score (Absolute) |  | 0.07 | 5.1 | NA | NA | 10.16 | 0.07 | NA | NA | 1.35 | 1.17 | 0.57 | 0.67 | 0.46 | 1.67 | 0.16 | 0.43 | 4.53 | 0.64 | 14.76 | 0.25 | 0 | 0.82 | 5.26 | NA | 5.62 | 0.27
Results from laboratories 1.05, 1.06, 1.09, 1.10 and 1.31 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 235b (Certified Value 128 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 154 | 130 | 160 | 127.89 | 118 | 127 | 113 | 110 | 129 | 124 | 126 | 125 | 109 | 133.5212 | 132 | 123.1 | 142.2 | 125 | 126 | 117 | 134.9 | 131 | 92.9351 | 153.3909 |  |  | 1
| 1 | 2 | 152 | 120 | 103 | 129.08 | 117 | 127 | 96 | 113 | 130 | 133 | 126 | 112 | 111 | 131.61 | 137 | 128 | 143.5 | 136 | 129 | 117 | 134.4 | 132 | 83.8745 | 153.3499 |  |  | 33
| 1 | 3 | 149 | 120 | 96 | 130.49 | 129 | 134 | 103 | 116 | 127 | 126 | 127 | 125 | 107 | 132.8363 | 136 | 117.1 | 142.1 | 132 | 119 | 119 | 136.6 | 134 | 82.9008 | 151.8766 |  |  | 16
| 1 | 4 | 153 | 120 | 92 | 132.6 | 128 | 135 | 115 | 113 | 131 | 138 | 127 | 125 | 104 | 134.386 | 135 | 100 | 141.5 | 133 | 118 | 118 | 133.8 | 136 | 80.5258 | 149.7069 |  |  | 128.2002
| 1 | 5 | 155 | 120 | 99 | 131.2233 | 122 | 129 | 113 | 108 | 125 | 135 | 125 | 124 | 106 | 132.5942 | 134 | 117.6 | 145.9 | 129 | 120 | 111 | 135.6 | 139 | 93.5408 | 151.8935 |  |  | 24
| 1 | 6 | 154 | 130 | 96 | 128.765 | 118 | 139 | 119 | 122 | 127 | 133 | 125 | 122 | 112 | 132.8869 | 135 | 125 | 142.2 | 124 | 120 | 114 | 136.1 | 136 | 90.9045 | 151.5221
| Mean |  | 152.8333 | 123.3333 | 107.6667 | 130.0081 | 122 | 131.8333 | 109.8333 | 113.6667 | 128.1667 | 131.5 | 126 | 122.1667 | 108.1667 | 132.9724 | 134.8333 | 118.4667 | 142.9 | 129.8333 | 122 | 116 | 135.2333 | 134.6667 | 87.4469 | 151.9566
| Median |  | 153.5 | 120 | 97.5 | 129.785 | 120 | 131.5 | 113 | 113 | 128 | 133 | 126 | 124.5 | 108 | 132.8616 | 135 | 120.35 | 142.2 | 130.5 | 120 | 117 | 135.25 | 135 | 87.3895 | 151.8851
| Std Dev. |  | 2.137 | 5.164 | 25.8972 | 1.7504 | 5.3292 | 4.916 | 8.5888 | 4.9261 | 2.2286 | 5.3944 | 0.89443 | 5.1153 | 3.0605 | 0.92998 | 1.7224 | 9.9867 | 1.6087 | 4.7081 | 4.4272 | 2.9665 | 1.0596 | 2.9439 | 5.6664 | 1.3608
| Rel.Std.Dev. |  | 1.398% | 4.187% | 24.05% | 1.346% | 4.368% | 3.729% | 7.82% | 4.334% | 1.739% | 4.102% | 0.7099% | 4.187% | 2.829% | 0.6994% | 1.277% | 8.43% | 1.126% | 3.626% | 3.629% | 2.557% | 0.7835% | 2.186% | 6.48% | 0.8955%
| PDM3 |  | 19.21% | -3.796% | -16.02% | 1.41% | -4.836% | 2.834% | -14.33% | -11.34% | -0.02614% | 2.574% | -1.716% | -4.706% | -15.63% | 3.723% | 5.174% | -7.592% | 11.47% | 1.274% | -4.836% | -9.517% | 5.486% | 5.044% | -31.79% | 18.53%
| Z-Score (Absolute) |  | 2.26 | 0.33 | 1.71 | 0.26 | 0.45 | 0.42 | 1.52 | 1.18 | 0.1 | 0.39 | 0.1 | 0.43 | 1.66 | 0.52 | 0.68 | 0.76 | 1.39 | 0.24 | 0.45 | 0.97 | 0.72 | 0.67 | 3.49 | 2.19
| Analytical results for Cs in OREAS 235b (Certified Value 9.98 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 10.17 | 8.7689 | 10 | 10.2 | 10.3 | 8.59 | 11.1 | 10.7 | 9.85 | 8.9 | 9.66 | 9.96 | 9.5 | 11.3 | 8.67 | 9 | 10.29 | 9.65 | 10.1 | 19.8419 | 9.8756 | [5 empty] | 1
| 1 | 2 | 9.96 | 8.7962 | 10.1 | 9.98 | 10.3 | 8.63 | 10.5 | 10.6 | 9.82 | 9.44 | 9.36 | 10.05 | 12.9 | 11.4 | 8.818 | 10 | 10.44 | 9.6 | 9.8 | 10.4301 | 9.707 | [5 empty] | 34
| 1 | 3 | 9.98 | 8.6669 | 9.7 | 10 | 10.2 | 9.11 | 10 | 10.4 | 9.59 | 9.03 | 9.28 | 10.9 | 11.8 | 11.3 | 8.874 | 10 | 10.11 | 9.3 | 10.2 | 10.2017 | 9.9153 | [5 empty] | 16
| 1 | 4 | 10.28 | 8.6062 | 9.8 | 10.2 | 10.3 | 9.79 | 10.1 | 9.6 | 9.99 | 10.7 | 9.53 | 10.35 | 12.7 | 11.1 | 8.779 | 9 | 10.2 | 9.39 | 9.8 | 10.1825 | 9.753 | [5 empty] | 9.9752
| 1 | 5 | 10.02 | 8.7648 | 9.9 | 10.5 | 10.2 | 10.55 | 9.7 | 9.5 | 9.78 | 10.05 | 10.15 | 10.35 | 9.3 | 11.3 | 8.788 | 10 | 10.27 | 9.16 | 9.8 | 10.4476 | 9.8486 | [5 empty] | 25
| 1 | 6 | 10.15 | 8.6506 | 10.1 | 9.77 | 10.4 | 10.23 | 10.8 | 9.5 | 9.61 | 9.96 | 10.1 | 10.35 | 9.7 | 11.1 | 8.972 | 9 | 9.75 | 9.64 | 9.7 | 10.286 | 9.6839
| Mean |  | 10.0933 | 8.7089 | 9.9333 | 10.1083 | 10.2833 | 9.4833 | 10.3667 | 10.05 | 9.7733 | 9.68 | 9.68 | 10.3267 | 10.9833 | 11.25 | 8.8168 | 9.5 | 10.1767 | 9.4567 | 9.9 | 11.8983 | 9.7973
| Median |  | 10.085 | 8.7158 | 9.95 | 10.1 | 10.3 | 9.45 | 10.3 | 10 | 9.8 | 9.7 | 9.595 | 10.35 | 10.75 | 11.3 | 8.803 | 9.5 | 10.235 | 9.495 | 9.8 | 10.3581 | 9.8008
| Std Dev. |  | 0.12644 | 0.077522 | 0.1633 | 0.25015 | 0.075277 | 0.83126 | 0.52789 | 0.57533 | 0.15188 | 0.68472 | 0.36949 | 0.32904 | 1.6714 | 0.12247 | 0.10116 | 0.54772 | 0.23577 | 0.20422 | 0.2 | 3.8931 | 0.095562
| Rel.Std.Dev. |  | 1.253% | 0.8901% | 1.644% | 2.475% | 0.732% | 8.765% | 5.092% | 5.725% | 1.554% | 7.074% | 3.817% | 3.186% | 15.22% | 1.089% | 1.147% | 5.766% | 2.317% | 2.16% | 2.02% | 32.72% | 0.9754%
| PDM3 |  | 1.184% | -12.69% | -0.42% | 1.334% | 3.089% | -4.931% | 3.924% | 0.7496% | -2.024% | -2.96% | -2.96% | 3.523% | 10.11% | 12.78% | -11.61% | -4.764% | 2.019% | -5.198% | -0.7541% | 19.28% | -1.784%
| Z-Score (Absolute) |  | 0.22 | 2.78 | 0.13 | 0.25 | 0.63 | 1.1 | 0.81 | 0.13 | 0.47 | 0.67 | 0.67 | 0.72 | 2.15 | 2.72 | 2.54 | NA | 0.4 | 1.16 | 0.2 | 4.13 | 0.42
Results from laboratory 1.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 235b (Certified Value 29.3 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 29.8 | 27.8829 | 30 | 28 | 27.7367 | 29 | 24.4 | 25 | 33.4 | 29.1 | 28.5 | 30.1 | 27.9 | 27.9 | 30.3977 | 27.3 | 27.01 | 30.1 | 33.03 | 30.3 | 29.9 | 30.5 | 30.18 | 30 | 2 | 28.8305 | 1
| 1 | 2 | 30 | 27.4237 | 30 | 28 | 27.91 | 29 | 25.6 | 28 | 30.1 | 28.3 | 27.8 | 29.1 | 28.8 | 28.2 | 30.6593 | 27 | 29.55 | 31.2 | 31.71 | 29.7 | 29.6 | 29.2 | 29.89 | 31 | 6 | 28.8074 | 35
| 1 | 3 | 30.1 | 27.0476 | 30 | 27 | 28.57 | 31 | 27.2 | 26 | 32.1 | 28 | 27.3 | 28.6 | 29 | 28.4 | 29.8097 | 27.6 | 27.7 | 30.7 | 32 | 29.8 | 28.4 | 29.1 | 30.87 | 30 | 6 | 28.8442 | 16
| 1 | 4 | 30.4 | 27.1493 | 28 | 27 | 28.5967 | 31 | 31.2 | 25 | 33.9 | 29.3 | 30.8 | 29.4 | 29.4 | 28.5 | 30.3747 | 27.4 | 28.89 | 31.7 | 31.81 | 29.6 | 28.6 | 28.3 | 29.91 | 31 | 6 | 28.7421 | 29.3305
| 1 | 5 | 30.3 | 25.7714 | 30 | 27 | 27.8833 | 30 | 32.4 | 25 | 32.7 | 28.4 | 28.8 | 30.7 | 29.7 | 29.9 | 30.388 | 27.5 | 28.3 | 32 | 32.53 | 29.8 | 28.5 | 28.9 | 30.54 | 29 | 5 | 28.6093 | 26
| 1 | 6 | 30.9 | 27.7447 | 30 | 27 | 29.27 | 30 | 32.2 | 26 | 32.1 | 29 | 29.4 | 30.2 | 29.1 | 29 | 30.146 | 27.4 | 29.01 | 31.6 | 31.79 | 29.4 | 29.7 | 29.7 | 30.8 | 33 | 5 | 28.5754
| Mean |  | 30.25 | 27.1699 | 29.6667 | 27.3333 | 28.3278 | 30 | 28.8333 | 25.8333 | 32.3833 | 28.6833 | 28.7667 | 29.6833 | 28.9833 | 28.65 | 30.2959 | 27.3667 | 28.41 | 31.2167 | 32.145 | 29.7667 | 29.1167 | 29.2833 | 30.365 | 30.6667 | 5 | 28.7348
| Median |  | 30.2 | 27.2865 | 30 | 27 | 28.24 | 30 | 29.2 | 25.5 | 32.4 | 28.7 | 28.65 | 29.75 | 29.05 | 28.45 | 30.3813 | 27.4 | 28.595 | 31.4 | 31.905 | 29.75 | 29.1 | 29.15 | 30.36 | 30.5 | 5.5 | 28.7747
| Std Dev. |  | 0.38341 | 0.75823 | 0.8165 | 0.5164 | 0.58996 | 0.89443 | 3.5336 | 1.169 | 1.3273 | 0.51929 | 1.2404 | 0.78337 | 0.61779 | 0.71204 | 0.28843 | 0.20656 | 0.93447 | 0.70828 | 0.52535 | 0.30111 | 0.68532 | 0.74944 | 0.43399 | 1.3663 | 1.5492 | 0.11627
| Rel.Std.Dev. |  | 1.267% | 2.791% | 2.752% | 1.889% | 2.083% | 2.981% | 12.26% | 4.525% | 4.099% | 1.81% | 4.312% | 2.639% | 2.132% | 2.485% | 0.952% | 0.7548% | 3.289% | 2.269% | 1.634% | 1.012% | 2.354% | 2.559% | 1.429% | 4.455% | 30.98% | 0.4046%
| PDM3 |  | 3.135% | -7.366% | 1.146% | -6.809% | -3.419% | 2.283% | -1.695% | -11.92% | 10.41% | -2.207% | -1.922% | 1.203% | -1.184% | -2.32% | 3.291% | -6.696% | -3.138% | 6.431% | 9.596% | 1.487% | -0.7291% | -0.1608% | 3.527% | 4.556% | -82.95% | -2.031%
| Z-Score (Absolute) |  | 0.97 | 1.52 | 0.5 | 1.38 | 0.58 | 0.77 | 0.17 | 2.59 | 2.69 | 0.3 | 0.23 | 0.51 | 0.05 | 0.32 | 1 | 1.36 | 0.52 | 1.75 | 2.5 | 0.58 | 0.05 | 0.19 | 1.06 | 1.3 | 19.4 | 0.25
| Analytical results for Dy in OREAS 235b (Certified Value 3.69 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS | 2
2
| 1 | 1 | 1.01: 3.1727; 1.03: 3.9; 1.05: 3.9; 1.06: 3.1; 1.07: 4.38; 1.08: 2.845; 1.09: 3.98; 1.10: 3.2; 1.11: 2.76; 1.12: 4.977 | 1
| 1 | 2 | 1.01: 3.0206; 1.03: 3.95; 1.05: 4.1; 1.06: 4.1; 1.07: 4.47; 1.08: 2.965; 1.09: 4.18; 1.10: 3.1; 1.11: 2.54; 1.12: 5.0306 | 9
| 1 | 3 | 1.01: 3.2154; 1.03: 3.85; 1.05: 3.8; 1.06: 4; 1.07: 4.34; 1.08: 3.349; 1.09: 3.83; 1.10: 3.2; 1.11: 2.77; 1.12: 5.2032 | 16
| 1 | 4 | 1.01: 3.1498; 1.03: 4.05; 1.05: 4; 1.06: 4.1; 1.07: 4.34; 1.08: 3.075; 1.09: 4.05; 1.10: 3; 1.11: 2.74; 1.12: 5.0423 | 3.6851
| 1 | 5 | 1.01: 3.2111; 1.03: 4.05; 1.05: 4.1; 1.06: 3; 1.07: 4.33; 1.08: 3.022; 1.09: 4.18; 1.10: 3; 1.11: 2.71; 1.12: 4.8772 | 27
| 1 | 6 | 1.01: 3.1492; 1.03: 3.9; 1.05: 4.1; 1.06: 3.1; 1.07: 4.2; 1.08: 2.969; 1.09: 3.89; 1.10: 3; 1.11: 2.74; 1.12: 4.9732
| Mean |  | 1.01: 3.1531; 1.03: 3.95; 1.05: 4; 1.06: 3.5667; 1.07: 4.3433; 1.08: 3.0375; 1.09: 4.0183; 1.10: 3.0833; 1.11: 2.71; 1.12: 5.0173
| Median |  | 1.01: 3.1613; 1.03: 3.925; 1.05: 4.05; 1.06: 3.55; 1.07: 4.34; 1.08: 2.9955; 1.09: 4.015; 1.10: 3.05; 1.11: 2.74; 1.12: 5.0038
| Std Dev. |  | 1.01: 0.071016; 1.03: 0.083666; 1.05: 0.12649; 1.06: 0.55015; 1.07: 0.087331; 1.08: 0.17069; 1.09: 0.14607; 1.10: 0.098319; 1.11: 0.08579; 1.12: 0.10822
| Rel.Std.Dev. |  | 1.01: 2.252%; 1.03: 2.118%; 1.05: 3.162%; 1.06: 15.42%; 1.07: 2.011%; 1.08: 5.62%; 1.09: 3.635%; 1.10: 3.189%; 1.11: 3.166%; 1.12: 2.157%
| PDM3 |  | 1.01: -14.44%; 1.03: 7.188%; 1.05: 8.544%; 1.06: -3.215%; 1.07: 17.86%; 1.08: -17.57%; 1.09: 9.042%; 1.10: -16.33%; 1.11: -26.46%; 1.12: 36.15%
| Z-Score (Absolute) |  | 1.01: 0.69; 1.03: 0.22; 1.05: 0.27; 1.06: 0.22; 1.07: 0.66; 1.08: 0.82; 1.09: 0.29; 1.10: 0.76; 1.11: 1.19; 1.12: 1.43
| Analytical results for Er in OREAS 235b (Certified Value 2.05 ppm) | [27 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.7214; 1.03: 2.2; 1.05: 2.1; 1.06: 1.8; 1.07: 2.41; 1.08: 1.556; 1.09: 2.24; 1.10: 1.6; 1.11: 1.4972; 1.12: 2.9212 | 1
| 1 | 2 | 1.01: 1.6852; 1.03: 2.25; 1.05: 2.2; 1.06: 2.8; 1.07: 2.39; 1.08: 1.606; 1.09: 2.25; 1.10: 1.6; 1.11: 1.4947; 1.12: 2.9857 | 10
| 1 | 3 | 1.01: 1.5744; 1.03: 2.2; 1.05: 2.2; 1.06: 3.1; 1.07: 2.33; 1.08: 1.684; 1.09: 2.15; 1.10: 1.6; 1.11: 1.5007; 1.12: 3.1031 | 16
| 1 | 4 | 1.01: 1.6404; 1.03: 2.1; 1.05: 2.2; 1.06: 2.6; 1.07: 2.34; 1.08: 1.664; 1.09: 2.17; 1.10: 1.5; 1.11: 1.4909; 1.12: 2.8918 | 2.0489
| 1 | 5 | 1.01: 1.6272; 1.03: 2.05; 1.05: 2.3; 1.06: 1.7; 1.07: 2.31; 1.08: 1.631; 1.09: 2.2; 1.10: 1.5; 1.11: 1.495; 1.12: 2.8909 | 28
| 1 | 6 | 1.01: 1.5707; 1.03: 2.15; 1.05: 2.3; 1.06: 1.8; 1.07: 2.28; 1.08: 1.58; 1.09: 2.36; 1.10: 1.6; 1.11: 1.4933; 1.12: 2.9269
| Mean |  | 1.01: 1.6365; 1.03: 2.1583; 1.05: 2.2167; 1.06: 2.3; 1.07: 2.3433; 1.08: 1.6202; 1.09: 2.2283; 1.10: 1.5667; 1.11: 1.4953; 1.12: 2.9533
| Median |  | 1.01: 1.6338; 1.03: 2.175; 1.05: 2.2; 1.06: 2.2; 1.07: 2.335; 1.08: 1.6185; 1.09: 2.22; 1.10: 1.6; 1.11: 1.4949; 1.12: 2.9241
| Std Dev. |  | 1.01: 0.059765; 1.03: 0.073598; 1.05: 0.075277; 1.06: 0.60663; 1.07: 0.048854; 1.08: 0.049081; 1.09: 0.075211; 1.10: 0.05164; 1.11: 0.0033628; 1.12: 0.081105
| Rel.Std.Dev. |  | 1.01: 3.652%; 1.03: 3.41%; 1.05: 3.396%; 1.06: 26.38%; 1.07: 2.085%; 1.08: 3.029%; 1.09: 3.375%; 1.10: 3.296%; 1.11: 0.2249%; 1.12: 2.746%
| PDM3 |  | 1.01: -20.12%; 1.03: 5.343%; 1.05: 8.19%; 1.06: 12.26%; 1.07: 14.37%; 1.08: -20.92%; 1.09: 8.759%; 1.10: -23.53%; 1.11: -27.02%; 1.12: 44.14%
| Z-Score (Absolute) |  | 1.01: 1.05; 1.03: 0.06; 1.05: 0.06; 1.06: 0.21; 1.07: 0.3; 1.08: 1.08; 1.09: 0.08; 1.10: 1.18; 1.11: 1.32; 1.12: 1.46
| Analytical results for Eu in OREAS 235b (Certified Value 1.29 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS | 2
3
| 1 | 1 | 1.01: 1.6211; 1.03: 1.3; 1.05: 1.25; 1.06: 1.3; 1.07: 1.38; 1.08: 1.291; 1.09: 1.39; 1.10: 1.3; 1.11: 2.3056; 1.12: 1.2455 | 1
| 1 | 2 | 1.01: 1.5927; 1.03: 1.3; 1.05: 1.27; 1.06: 1.5; 1.07: 1.42; 1.08: 1.329; 1.09: 1.43; 1.10: 1.2; 1.11: 1.1977; 1.12: 1.2399 | 11
| 1 | 3 | 1.01: 1.6309; 1.03: 1.25; 1.05: 1.22; 1.06: 1.3; 1.07: 1.4; 1.08: 1.32; 1.09: 1.3; 1.10: 1.2; 1.11: 1.1613; 1.12: 1.2752 | 16
| 1 | 4 | 1.01: 1.628; 1.03: 1.25; 1.05: 1.27; 1.06: 1.5; 1.07: 1.38; 1.08: 1.349; 1.09: 1.67; 1.10: 1.2; 1.11: 1.1537; 1.12: 1.2261 | 1.2891
| 1 | 5 | 1.01: 1.6482; 1.03: 1.35; 1.05: 1.27; 1.06: 1.2; 1.07: 1.39; 1.08: 1.337; 1.09: 1.34; 1.10: 1.2; 1.11: 1.1892; 1.12: 1.2108 | 29
| 1 | 6 | 1.01: 1.6638; 1.03: 1.3; 1.05: 1.22; 1.06: 1.2; 1.07: 1.4; 1.08: 1.324; 1.09: 1.31; 1.10: 1.2; 1.11: 1.153; 1.12: 1.2268
| Mean |  | 1.01: 1.6308; 1.03: 1.2917; 1.05: 1.25; 1.06: 1.3333; 1.07: 1.395; 1.08: 1.325; 1.09: 1.4067; 1.10: 1.2167; 1.11: 1.3601; 1.12: 1.2374
| Median |  | 1.01: 1.6294; 1.03: 1.3; 1.05: 1.26; 1.06: 1.3; 1.07: 1.395; 1.08: 1.3265; 1.09: 1.365; 1.10: 1.2; 1.11: 1.1752; 1.12: 1.2334
| Std Dev. |  | 1.01: 0.024298; 1.03: 0.037639; 1.05: 0.024495; 1.06: 0.13663; 1.07: 0.015166; 1.08: 0.019586; 1.09: 0.13808; 1.10: 0.040825; 1.11: 0.46359; 1.12: 0.02212
| Rel.Std.Dev. |  | 1.01: 1.49%; 1.03: 2.914%; 1.05: 1.96%; 1.06: 10.25%; 1.07: 1.087%; 1.08: 1.478%; 1.09: 9.816%; 1.10: 3.355%; 1.11: 34.09%; 1.12: 1.788%
| PDM3 |  | 1.01: 26.5%; 1.03: 0.1955%; 1.05: -3.037%; 1.06: 3.428%; 1.07: 8.211%; 1.08: 2.781%; 1.09: 9.116%; 1.10: -5.622%; 1.11: 5.503%; 1.12: -4.016%
| Z-Score (Absolute) |  | 1.01: 3.33; 1.03: 0.59; 1.05: 1.07; 1.06: NA; 1.07: 0.61; 1.08: 0.2; 1.09: 0.74; 1.10: NA; 1.11: 0.2; 1.12: 1.22
Results from laboratories 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 235b (Certified Value 4.15 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 4.26 | 4.0892 | 4.23 | 3.99 | 4.2767 | 4.56 | 3.88 | 4.3 | 4.03 | 4.13 | 3.96 | 4.07 | 3.95 | 4.0902 | 4.2588 | 4.092 | 4.16 | 4.33 | 4.23 | 4.24 | 4.3 | 4.131 | 4.2436 | 3.8674 |  |  | 1
| 1 | 2 | 4.19 | 4.1373 | 4.28 | 3.81 | 4.3033 | 4.53 | 3.83 | 4.14 | 4.01 | 4.12 | 4.31 | 4.12 | 3.96 | 4.1065 | 4.2389 | 4.064 | 4.18 | 4.37 | 4.3 | 4.32 | 4.17 | 4.1094 | 4.2637 | 4.2076 |  |  | 0
| 1 | 3 | 4.16 | 4.0956 | 4.16 | 3.98 | 4.31 | 4.51 | 3.85 | 4.66 | 4.25 | 4.03 | 4.11 | 4.05 | 4.04 | 4.1137 | 4.2221 | 4.106 | 4.17 | 4.32 | 4.22 | 4.27 | 4.15 | 4.1833 | 4.1395 | 4.1728 |  |  | 16
| 1 | 4 | 4.25 | 4.0454 | 4.13 | 3.97 | 4.295 | 4.57 | 4.07 | 4.25 | 4.02 | 4.12 | 4.38 | 4.08 | 4.02 | 4.1545 | 4.2291 | 4.057 | 4.1 | 4.37 | 4.29 | 4.28 | 4.17 | 4.0957 | 4.1622 | 4.1879 |  |  | 4.1515
| 1 | 5 | 4.25 | 4.0197 | 4.26 | 3.88 | 4.2833 | 4.51 | 4.14 | 4.24 | 3.99 | 4 | 4.35 | 4.19 | 4.04 | 4.1575 | 4.2104 | 4.022 | 4.14 | 4.33 | 4.25 | 4.27 | 4.07 | 4.1367 | 4.0839 | 4.1678 |  |  | 30
| 1 | 6 | 4.25 | 4.1089 | 4.28 | 3.78 | 4.295 | 4.52 | 4.11 | 4.12 | 4.03 | 4.02 | 4.26 | 4.18 | 4.01 | 4.2289 | 4.2495 | 4.078 | 4.15 | 4.33 | 4.29 | 4.32 | 4.14 | 4.1565 | 4.0877 | 4.1855
| Mean |  | 4.2267 | 4.0827 | 4.2233 | 3.9017 | 4.2939 | 4.5333 | 3.98 | 4.285 | 4.055 | 4.07 | 4.2283 | 4.115 | 4.0033 | 4.1419 | 4.2348 | 4.0698 | 4.15 | 4.3417 | 4.2633 | 4.2833 | 4.1667 | 4.1354 | 4.1634 | 4.1315
| Median |  | 4.25 | 4.0924 | 4.245 | 3.925 | 4.295 | 4.525 | 3.975 | 4.245 | 4.025 | 4.075 | 4.285 | 4.1 | 4.015 | 4.1341 | 4.234 | 4.071 | 4.155 | 4.33 | 4.27 | 4.275 | 4.16 | 4.1338 | 4.1509 | 4.1791
| Std Dev. |  | 0.041312 | 0.042978 | 0.064083 | 0.09196 | 0.012323 | 0.02582 | 0.14142 | 0.19614 | 0.096695 | 0.05933 | 0.16216 | 0.058907 | 0.039328 | 0.050383 | 0.017862 | 0.029505 | 0.028284 | 0.022286 | 0.034448 | 0.031411 | 0.075011 | 0.031663 | 0.076309 | 0.13012
| Rel.Std.Dev. |  | 0.9774% | 1.053% | 1.517% | 2.357% | 0.287% | 0.5696% | 3.553% | 4.577% | 2.385% | 1.458% | 3.835% | 1.432% | 0.9824% | 1.216% | 0.4218% | 0.725% | 0.6815% | 0.5133% | 0.808% | 0.7333% | 1.8% | 0.7657% | 1.833% | 3.15%
| PDM3 |  | 1.81% | -1.658% | 1.729% | -6.019% | 3.429% | 9.196% | -4.132% | 3.215% | -2.325% | -1.964% | 1.85% | -0.8801% | -3.57% | -0.2326% | 2.005% | -1.968% | -0.03706% | 4.58% | 2.693% | 3.175% | 0.3644% | -0.3879% | 0.2865% | -0.4827%
| Z-Score (Absolute) |  | 0.57 | 0.61 | 0.55 | 2.09 | 1.12 | 3.08 | 1.45 | 1.05 | 0.83 | 0.71 | 0.59 | 0.34 | 1.26 | 0.12 | 0.64 | 0.71 | 0.05 | 1.51 | 0.87 | 1.04 | 0.08 | 0.17 | 0.05 | 0.21
| Analytical results for Ga in OREAS 235b (Certified Value 19.9 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 20 | 23.8425 | 19.4 | 21.5 | 21.6733 | 19.7 | 17.62 | 22.8 | 22 | 19.45 | 19.4 | 20.6 | 18.95 | 19.4 | 18.9652 | 18 | 18.6 | 19.16 | 21.28 | 18.91 | 40.4457 | 24 | 19.4131 |  |  |  | 1
| 1 | 2 | 20.62 | 24.453 | 19.8 | 21.8 | 21.89 | 19.7 | 17.49 | 22.5 | 21 | 19.3 | 20.7 | 19.7 | 19.6 | 19.3 | 19.2085 | 18.4 | 19.9 | 19.15 | 19.18 | 18.93 | 22.2739 | 23 | 19.5281 |  |  |  | 39
| 1 | 3 | 19.94 | 24.035 | 20.2 | 21.3 | 21.26 | 19.7 | 18.55 | 19.7 | 20 | 18.65 | 19.15 | 19.1 | 19.6 | 19.77 | 19.142 | 17.8 | 20.5 | 18.37 | 19.87 | 18.71 | 21.5759 | 21 | 19.9959 |  |  |  | 16
| 1 | 4 | 20.47 | 25.0705 | 19.8 | 22 | 22.3833 | 19.9 | 20.7 | 20.8 | 20 | 19.8 | 20.4 | 19.6 | 19.75 | 19.21 | 18.9908 | 17.9 | 19.5 | 19.32 | 18.92 | 18.46 | 20.4389 | 24 | 19.5294 |  |  |  | 19.8956
| 1 | 5 | 19.95 | 23.313 | 19.6 | 21.2 | 22.27 | 19.9 | 20.07 | 20.5 | 21 | 19.15 | 21 | 20.5 | 19.8 | 18.8 | 18.9539 | 17.7 | 19.7 | 18.42 | 18.98 | 18.11 | 22.3861 | 24 | 19.6165 |  |  |  | 31
| 1 | 6 | 20.07 | 24.301 | 20 | 20.7 | 22.365 | 20 | 20.75 | 22 | 20 | 19.2 | 21.1 | 20.5 | 19.65 | 19.63 | 19.1772 | 17.8 | 18.9 | 17.91 | 21.01 | 18.39 | 21.3616 | 22 | 19.3946
| Mean |  | 20.175 | 24.1692 | 19.8 | 21.4167 | 21.9736 | 19.8167 | 19.1967 | 21.3833 | 20.6667 | 19.2583 | 20.2917 | 20 | 19.5583 | 19.3517 | 19.0729 | 17.9333 | 19.5167 | 18.7217 | 19.8733 | 18.585 | 24.747 | 23 | 19.5796
| Median |  | 20.035 | 24.168 | 19.8 | 21.4 | 22.08 | 19.8 | 19.31 | 21.4 | 20.5 | 19.25 | 20.55 | 20.1 | 19.625 | 19.35 | 19.0664 | 17.85 | 19.6 | 18.785 | 19.525 | 18.585 | 21.9249 | 23.5 | 19.5288
| Std Dev. |  | 0.29413 | 0.59466 | 0.28284 | 0.46224 | 0.45062 | 0.13292 | 1.5002 | 1.2319 | 0.8165 | 0.3787 | 0.8285 | 0.61968 | 0.3089 | 0.34114 | 0.11538 | 0.25033 | 0.68823 | 0.56694 | 1.0448 | 0.32235 | 7.7229 | 1.2649 | 0.21991
| Rel.Std.Dev. |  | 1.458% | 2.46% | 1.428% | 2.158% | 2.051% | 0.6707% | 7.815% | 5.761% | 3.951% | 1.966% | 4.083% | 3.098% | 1.579% | 1.763% | 0.6049% | 1.396% | 3.526% | 3.028% | 5.257% | 1.734% | 31.21% | 5.5% | 1.123%
| PDM3 |  | 1.405% | 21.48% | -0.4803% | 7.645% | 10.44% | -0.3965% | -3.513% | 7.478% | 3.876% | -3.203% | 1.991% | 0.5249% | -1.695% | -2.734% | -4.135% | -9.863% | -1.904% | -5.9% | -0.1117% | -6.587% | 24.38% | 15.6% | -1.588%
| Z-Score (Absolute) |  | 0.39 | 4.73 | 0.02 | 1.74 | 2.35 | 0 | 0.67 | 1.7 | 0.92 | 0.61 | 0.52 | 0.2 | 0.28 | 0.51 | 0.81 | 2.05 | 0.33 | 1.19 | 0.06 | 1.34 | 5.36 | 3.46 | 0.26
| Analytical results for Gd in OREAS 235b (Certified Value 5.23 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 6.9185; 1.03: 5.6; 1.05: 5.1; 1.06: 4.9; 1.07: 5.3; 1.08: 5.077; 1.09: 5.56; 1.10: 4.7; 1.11: 10.0813; 1.12: 0.82149 |  |  | 1
| 1 | 2 | 1.01: 7.1097; 1.03: 5.6; 1.05: 5.3; 1.06: 5.7; 1.07: 5.4; 1.08: 5.273; 1.09: 5.57; 1.10: 4.5; 1.11: 5.0791; 1.12: 0.85078 |  |  | 12
| 1 | 3 | 1.01: 7.1281; 1.03: 5.6; 1.05: 5.3; 1.06: 5.8; 1.07: 5.2; 1.08: 5.252; 1.09: 5.37; 1.10: 4.7; 1.11: 5.1283; 1.12: 0.85934 |  |  | 16
| 1 | 4 | 1.01: 7.1543; 1.03: 5.4; 1.05: 5.4; 1.06: 6.3; 1.07: 5.2; 1.08: 5.537; 1.09: 5.67; 1.10: 4.5; 1.11: 5.0738; 1.12: 0.83036 |  |  | 5.2277
| 1 | 5 | 1.01: 7.138; 1.03: 5.4; 1.05: 5.3; 1.06: 4.5; 1.07: 5.2; 1.08: 5.357; 1.09: 5.19; 1.10: 4.5; 1.11: 5.2002; 1.12: 0.82472 |  |  | 32
| 1 | 6 | 1.01: 7.1061; 1.03: 5.6; 1.05: 5.1; 1.06: 4.8; 1.07: 5.3; 1.08: 5.284; 1.09: 5.19; 1.10: 4.6; 1.11: 5.1853; 1.12: 0.83147
| Mean |  | 1.01: 7.0925; 1.03: 5.5333; 1.05: 5.25; 1.06: 5.3333; 1.07: 5.2667; 1.08: 5.2967; 1.09: 5.425; 1.10: 4.5833; 1.11: 5.958; 1.12: 0.83636
| Median |  | 1.01: 7.1189; 1.03: 5.6; 1.05: 5.3; 1.06: 5.3; 1.07: 5.25; 1.08: 5.2785; 1.09: 5.465; 1.10: 4.55; 1.11: 5.1568; 1.12: 0.83091
| Std Dev. |  | 1.01: 0.087076; 1.03: 0.10328; 1.05: 0.12247; 1.06: 0.70048; 1.07: 0.08165; 1.08: 0.14991; 1.09: 0.20628; 1.10: 0.098319; 1.11: 2.0207; 1.12: 0.015182
| Rel.Std.Dev. |  | 1.01: 1.228%; 1.03: 1.866%; 1.05: 2.333%; 1.06: 13.13%; 1.07: 1.55%; 1.08: 2.83%; 1.09: 3.802%; 1.10: 2.145%; 1.11: 33.92%; 1.12: 1.815%
| PDM3 |  | 1.01: 35.67%; 1.03: 5.846%; 1.05: 0.4264%; 1.06: 2.02%; 1.07: 0.7452%; 1.08: 1.319%; 1.09: 3.774%; 1.10: -12.33%; 1.11: 13.97%; 1.12: -84%
| Z-Score (Absolute) |  | 1.01: 7.3; 1.03: 0.9; 1.05: 0.27; 1.06: 0.08; 1.07: 0.2; 1.08: 0.08; 1.09: 0.45; 1.10: 3.01; 1.11: 2.64; 1.12: 18.4
| Analytical results for Ge in OREAS 235b (Indicative Value 0.21 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.06; 1.05: 1.08; 1.06: 1.09; 1.07: 1.10; 1.08: 1.11; 1.09: 1.12; 1.10: 1.13; 1.11: 1.14; 1.12: 1.15; 1.13: 1.16; 1.14: 1.26; 1.15: 1.27; 1.16: 1.30 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*OES/MS; 1.05: 4A*OES; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*OES/MS; 1.11: 4A*OES/MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*MS; 1.15: 4A*OES/MS; 1.16: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 1.7; 1.03: < 0.1; 1.05: 1.1; 1.06: 0.18; 1.07: 1.7; 1.08: 4.1; 1.09: 0.18; 1.10: 0.24; 1.11: 0.1; 1.12: 0.16; 1.13: 0.46; 1.14: < 0.05; 1.15: 0.13; 1.16: 0.4684 |  |  | 1
| 1 | 2 | 1.01: 1.6; 1.03: 0.1; 1.05: 1.4; 1.06: 0.17; 1.07: 1.7; 1.08: 4.6; 1.09: 0.18; 1.10: 0.25; 1.11: 0.09; 1.12: 0.18; 1.13: 0.41; 1.14: < 0.05; 1.15: 0.12; 1.16: 0.3747 |  |  | 6
| 1 | 3 | 1.01: 1.8; 1.03: < 0.1; 1.05: 1.4; 1.06: 0.17; 1.07: 1.6; 1.08: 4.1; 1.09: 0.16; 1.10: 0.21; 1.11: 0.08; 1.12: 0.16; 1.13: 0.49; 1.14: < 0.05; 1.15: 0.13; 1.16: 0.2471 |  |  | 16
| 1 | 4 | 1.01: 1.8; 1.03: 0.1; 1.05: 1.2; 1.06: 0.19; 1.07: 2.1; 1.08: 4.1; 1.09: 0.19; 1.10: 0.16; 1.11: 0.09; 1.12: 0.17; 1.13: 0.27; 1.14: < 0.05; 1.15: 0.13; 1.16: 0.2833 |  |  | 0.20666
| 1 | 5 | 1.01: 1.8; 1.03: 0.2; 1.05: 1.3; 1.06: 0.19; 1.07: 2; 1.08: 4.1; 1.09: 0.19; 1.10: 0.25; 1.11: 0.09; 1.12: 0.18; 1.13: 0.38; 1.14: < 0.05; 1.15: 0.12; 1.16: 0.2914 |  |  | 12
| 1 | 6 | 1.01: 1.8; 1.03: < 0.1; 1.05: 1.2; 1.06: 0.19; 1.07: 1.8; 1.08: 4.1; 1.09: 0.18; 1.10: 0.24; 1.11: 0.11; 1.12: 0.19; 1.13: 0.58; 1.14: < 0.05; 1.15: 0.14; 1.16: 0.315
| Mean |  | 1.01: 1.75; 1.03: 0.13333; 1.05: 1.2667; 1.06: 0.18167; 1.07: 1.8167; 1.08: 4.1833; 1.09: 0.18; 1.10: 0.225; 1.11: 0.093333; 1.12: 0.17333; 1.13: 0.43167; 1.15: 0.12833; 1.16: 0.32998
| Median |  | 1.01: 1.8; 1.03: 0.1; 1.05: 1.25; 1.06: 0.185; 1.07: 1.75; 1.08: 4.1; 1.09: 0.18; 1.10: 0.24; 1.11: 0.09; 1.12: 0.175; 1.13: 0.435; 1.15: 0.13; 1.16: 0.3032
| Std Dev. |  | 1.01: 0.083666; 1.03: 0.057735; 1.05: 0.12111; 1.06: 0.0098319; 1.07: 0.19408; 1.08: 0.20412; 1.09: 0.010954; 1.10: 0.035071; 1.11: 0.010328; 1.12: 0.012111; 1.13: 0.10534; 1.15: 0.0075277; 1.16: 0.079901
| Rel.Std.Dev. |  | 1.01: 4.781%; 1.03: 43.3%; 1.05: 9.561%; 1.06: 5.412%; 1.07: 10.68%; 1.08: 4.879%; 1.09: 6.086%; 1.10: 15.59%; 1.11: 11.07%; 1.12: 6.987%; 1.13: 24.4%; 1.15: 5.866%; 1.16: 24.21%
| PDM3 |  | 1.01: 746.8%; 1.03: -35.48%; 1.05: 512.9%; 1.06: -12.1%; 1.07: 779%; 1.08: 1924%; 1.09: -12.9%; 1.10: 8.872%; 1.11: -54.84%; 1.12: -16.13%; 1.13: 108.9%; 1.15: -37.9%; 1.16: 59.67%
| Z-Score (Absolute) |  | 1.01: 8.75; 1.03: 0.63; 1.05: 6.02; 1.06: 0.12; 1.07: 9.13; 1.08: 22.52; 1.09: 0.13; 1.10: 0.12; 1.11: 0.62; 1.12: 0.17; 1.13: 1.29; 1.14: 1.01; 1.15: 0.42; 1.16: 0.72
| Analytical results for Hf in OREAS 235b (Certified Value 3.92 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 4.17 | 3.9629 | 7.6 | 4.4 | 3.8 | 3.5 | 4.5 | 4.2 | 4 | 3.8 | 3.9 | 3.8 | 4.24 | 4.5 | 3.82 | 4.24 | 4 | 3.62 | 6.6584 | 2.8045 | 6.2132 | [5 empty] | 1
| 1 | 2 | 3.98 | 4.1378 | 7.6 | 3.5 | 3.9 | 3.6 | 4.5 | 4.5 | 4.1 | 3.7 | 3.8 | 3.7 | 4.26 | 4.5 | 3.81 | 4.39 | 3.8 | 3.45 | 3.4617 | 2.8461 | 6.1679 | [5 empty] | 25
| 1 | 3 | 4.23 | 4.102 | 8.2 | 3 | 3.9 | 3.6 | 3.9 | 4.4 | 4 | 3.7 | 4 | 3.8 | 3.99 | 4.4 | 3.89 | 4.18 | 3.7 | 3.58 | 3.4221 | 2.8045 | 6.4272 | [5 empty] | 16
| 1 | 4 | 4.18 | 4.0718 | 7.8 | 2.8 | 4.1 | 3.9 | 4.5 | 3.9 | 4.2 | 4.2 | 4.1 | 3.8 | 4.29 | 4.4 | 3.68 | 4.26 | 3.8 | 3.44 | 3.3773 | 2.7546 | 6.164 | [5 empty] | 3.9194
| 1 | 5 | 4.33 | 4.0215 | 8 | 3 | 4.1 | 3.8 | 3.9 | 4.4 | 4 | 4.2 | 3.9 | 3.8 | 3.7 | 4.4 | 3.62 | 4.27 | 3.6 | 3.42 | 3.4267 | 2.555 | 6.1425 | [5 empty] | 33
| 1 | 6 | 4.25 | 4.1003 | 8.4 | 2.9 | 4 | 3.9 | 4 | 4.6 | 4 | 4.1 | 4.1 | 3.8 | 4.1 | 4.3 | 3.67 | 4.15 | 3.9 | 3.4 | 3.3829 | 2.8461 | 6.234
| Mean |  | 4.19 | 4.066 | 7.9333 | 3.2667 | 3.9667 | 3.7167 | 4.2167 | 4.3333 | 4.05 | 3.95 | 3.9667 | 3.7833 | 4.0967 | 4.4167 | 3.7483 | 4.2483 | 3.8 | 3.485 | 3.9548 | 2.7685 | 6.2248
| Median |  | 4.205 | 4.086 | 7.9 | 3 | 3.95 | 3.7 | 4.25 | 4.4 | 4 | 3.95 | 3.95 | 3.8 | 4.17 | 4.4 | 3.745 | 4.25 | 3.8 | 3.445 | 3.4244 | 2.8045 | 6.1905
| Std Dev. |  | 0.11781 | 0.063669 | 0.3266 | 0.60553 | 0.12111 | 0.17224 | 0.31252 | 0.25033 | 0.083666 | 0.2429 | 0.12111 | 0.040825 | 0.22509 | 0.075277 | 0.1061 | 0.083766 | 0.14142 | 0.091597 | 1.3248 | 0.10992 | 0.10481
| Rel.Std.Dev. |  | 2.812% | 1.566% | 4.117% | 18.54% | 3.053% | 4.634% | 7.411% | 5.777% | 2.066% | 6.149% | 3.053% | 1.079% | 5.495% | 1.704% | 2.831% | 1.972% | 3.722% | 2.628% | 33.5% | 3.97% | 1.684%
| PDM3 |  | 6.905% | 3.743% | 102.4% | -16.65% | 1.207% | -5.172% | 7.586% | 10.56% | 3.333% | 0.7817% | 1.207% | -3.471% | 4.524% | 12.69% | -4.364% | 8.393% | -3.045% | -11.08% | 0.9055% | -29.36% | 58.82%
| Z-Score (Absolute) |  | 0.67 | 0.3 | 11.98 | 2.11 | 0 | 0.75 | 0.75 | 1.11 | 0.25 | 0.05 | 0 | 0.55 | 0.39 | 1.36 | 0.66 | 0.85 | 0.5 | 1.45 | 0.04 | 3.62 | 6.82
| Analytical results for Hg in OREAS 235b (Indicative Value 0.025 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.30 |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.03: 4A*MS |  |  | 3
3
| 1 | 1 | 1.01: < 2; 1.03: 0.027 |  |  | 1
| 1 | 2 | 1.01: < 2; 1.03: 0.0314 |  |  | 13
| 1 | 3 | 1.01: < 2; 1.03: 0.0209 |  |  | 16
| 1 | 4 | 1.01: < 2; 1.03: 0.0287 |  |  | 0.025383
| 1 | 5 | 1.01: < 2; 1.03: 0.0229 |  |  | 9
| 1 | 6 | 1.01: < 2; 1.03: 0.0214
| Mean |  | 1.03: 0.025383
| Median |  | 1.03: 0.02495
| Std Dev. |  | 1.03: 0.0042883
| Rel.Std.Dev. |  | 1.03: 16.89%
| PDM3 |  | 1.03: 1.332e-13%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for Ho in OREAS 235b (Certified Value 0.71 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.55446; 1.03: 0.76; 1.05: 0.8; 1.06: 0.6; 1.07: 0.77; 1.08: 0.534; 1.09: 0.79; 1.10: 0.6; 1.11: 0.5491; 1.12: 1.0134 |  |  | 1
| 1 | 2 | 1.01: 0.57719; 1.03: 0.76; 1.05: 0.8; 1.06: 1; 1.07: 0.77; 1.08: 0.559; 1.09: 0.8; 1.10: 0.6; 1.11: 0.5279; 1.12: 1.0205 |  |  | 14
| 1 | 3 | 1.01: 0.59132; 1.03: 0.72; 1.05: 0.8; 1.06: 0.7; 1.07: 0.75; 1.08: 0.603; 1.09: 0.74; 1.10: 0.6; 1.11: 0.5779; 1.12: 1.0453 |  |  | 16
| 1 | 4 | 1.01: 0.55598; 1.03: 0.72; 1.05: 0.8; 1.06: 0.8; 1.07: 0.8; 1.08: 0.558; 1.09: 0.77; 1.10: 0.6; 1.11: 0.5241; 1.12: 1.0086 |  |  | 0.71004
| 1 | 5 | 1.01: 0.56253; 1.03: 0.8; 1.05: 0.8; 1.06: 0.6; 1.07: 0.75; 1.08: 0.591; 1.09: 0.77; 1.10: 0.6; 1.11: 0.524; 1.12: 0.98844 |  |  | 34
| 1 | 6 | 1.01: 0.56362; 1.03: 0.74; 1.05: 0.8; 1.06: 0.6; 1.07: 0.76; 1.08: 0.569; 1.09: 0.77; 1.10: 0.6; 1.11: 0.5235; 1.12: 1.008
| Mean |  | 1.01: 0.56752; 1.03: 0.75; 1.05: 0.8; 1.06: 0.71667; 1.07: 0.76667; 1.08: 0.569; 1.09: 0.77333; 1.10: 0.6; 1.11: 0.53775; 1.12: 1.014
| Median |  | 1.01: 0.56308; 1.03: 0.75; 1.05: 0.8; 1.06: 0.65; 1.07: 0.765; 1.08: 0.564; 1.09: 0.77; 1.10: 0.6; 1.11: 0.526; 1.12: 1.011
| Std Dev. |  | 1.01: 0.014167; 1.03: 0.030332; 1.05: 1.2162e-16; 1.06: 0.16021; 1.07: 0.018619; 1.08: 0.024844; 1.09: 0.020656; 1.10: 0; 1.11: 0.021983; 1.12: 0.018653
| Rel.Std.Dev. |  | 1.01: 2.496%; 1.03: 4.044%; 1.05: 1.52e-14%; 1.06: 22.35%; 1.07: 2.429%; 1.08: 4.366%; 1.09: 2.671%; 1.10: 0%; 1.11: 4.088%; 1.12: 1.84%
| PDM3 |  | 1.01: -20.07%; 1.03: 5.628%; 1.05: 12.67%; 1.06: 0.9336%; 1.07: 7.975%; 1.08: -19.86%; 1.09: 8.914%; 1.10: -15.5%; 1.11: -24.26%; 1.12: 42.81%
| Z-Score (Absolute) |  | 1.01: 0.68; 1.03: 0; 1.05: NA; 1.06: NA; 1.07: 0.06; 1.08: 0.67; 1.09: 0.09; 1.10: NA; 1.11: 0.79; 1.12: 0.98
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 235b (Certified Value 0.07 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | 0.07 | 0.1 | < 0.1 | < 2 | 0.07 | 0.044 | 0.08 | 0.08 | 0.073 | 0.069 | 0.075 | 0.066 | 0.08 | 0.06 | 0.082 | 0.067 | 0.07 | 0.0857 | 0.06922 | [7 empty] | 1
| 1 | 2 | 0.07 | 0.1 | < 0.1 | < 2 | 0.07 | 0.04 | 0.1 | 0.07 | 0.074 | 0.071 | 0.065 | 0.066 | 0.07 | 0.07 | 0.067 | 0.062 | 0.07 | 0.0889 | 0.06796 | [7 empty] | 27
| 1 | 3 | 0.07 | 0.1 | < 0.1 | < 2 | 0.07 | 0.04 | 0.08 | 0.08 | 0.063 | 0.067 | 0.066 | 0.069 | 0.07 | 0.07 | 0.093 | 0.064 | 0.07 | 0.0832 | 0.06993 | [7 empty] | 16
| 1 | 4 | 0.07 | 0.1 | < 0.1 | < 2 | 0.07 | 0.045 | 0.1 | 0.08 | 0.069 | 0.08 | 0.07 | 0.069 | 0.08 | 0.06 | 0.081 | 0.059 | 0.06 | 0.0861 | 0.06816 | [7 empty] | 0.069702
| 1 | 5 | 0.07 | 0.1 | < 0.1 | < 2 | 0.07 | 0.045 | 0.1 | 0.07 | 0.066 | 0.071 | 0.071 | 0.068 | 0.07 | 0.07 | 0.082 | 0.06 | 0.06 | 0.0861 | 0.06993 | [7 empty] | 35
| 1 | 6 | 0.07 | 0.05 | < 0.1 | < 2 | 0.07 | 0.045 | 0.12 | 0.07 | 0.07 | 0.068 | 0.072 | 0.071 | 0.07 | 0.07 | 0.074 | 0.069 | 0.06 | 0.0889 | 0.06912
| Mean |  | 0.07 | 0.091667 |  |  | 0.07 | 0.043167 | 0.096667 | 0.075 | 0.069167 | 0.071 | 0.069833 | 0.068167 | 0.073333 | 0.066667 | 0.079833 | 0.0635 | 0.065 | 0.086483 | 0.069053
| Median |  | 0.07 | 0.1 |  |  | 0.07 | 0.0445 | 0.1 | 0.075 | 0.0695 | 0.07 | 0.0705 | 0.0685 | 0.07 | 0.07 | 0.0815 | 0.063 | 0.065 | 0.0861 | 0.06917
| Std Dev. |  | 0 | 0.020412 |  |  | 0 | 0.0024833 | 0.015055 | 0.0054772 | 0.0041673 | 0.0046904 | 0.0037639 | 0.0019408 | 0.005164 | 0.005164 | 0.0087502 | 0.003937 | 0.0054772 | 0.0021619 | 0.00084412
| Rel.Std.Dev. |  | 0% | 22.27% |  |  | 0% | 5.753% | 15.57% | 7.303% | 6.025% | 6.606% | 5.39% | 2.847% | 7.042% | 7.746% | 10.96% | 6.2% | 8.427% | 2.5% | 1.222%
| PDM3 |  | 0.4282% | 31.51% |  |  | 0.4282% | -38.07% | 38.69% | 7.602% | -0.7674% | 1.863% | 0.1891% | -2.202% | 5.21% | -4.354% | 14.54% | -8.897% | -6.745% | 24.08% | -0.93%
| Z-Score (Absolute) |  | 0 | 2.92 | 2.7 | 125.42 | 0 | 3.62 | 3.6 | 0.67 | 0.11 | 0.13 | 0.02 | 0.25 | 0.45 | 0.45 | 1.33 | 0.88 | 0.67 | 2.22 | 0.13
| Analytical results for Ir in OREAS 235b (Indicative Value 0.002 ppm) | [27 empty] | Indicative
Ir | Round | Replicate | 1.01: Lab | [7 empty] | 1
| No | No | 1.01: 1.30 | [7 empty] | ppm
|  |  | 1.01: 4A*MS | [7 empty] | 3
3
| 1 | 1 | 1.01: 0.002 | [7 empty] | 1
| 1 | 2 | 1.01: 0.0027 | [7 empty] | 4
| 1 | 3 | 1.01: 0.0031 | [7 empty] | 16
| 1 | 4 | 1.01: 0.0027 | [7 empty] | 0.0023667
| 1 | 5 | 1.01: 0.0025 | [7 empty] | 10
| 1 | 6 | 1.01: 0.0012
| Mean |  | 1.01: 0.0023667
| Median |  | 1.01: 0.0026
| Std Dev. |  | 1.01: 0.00067429
| Rel.Std.Dev. |  | 1.01: 28.49%
| PDM3 |  | 1.01: -1.665e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for K in OREAS 235b (Certified Value 2.76 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 2.892 | 2.811 | 2.68 | 1.62 | 2.3424 | 2.866 | 2.68 | 2.85 | 2.79 | 2.77 | 2.62 | 2.71 | 2.68 | 2.7397 | 2.7243 | 2.781 | 2.69 | 2.72 | 2.42 | 2.79 | 2.73 | 2.7144 | 3.0783 | 3.1532 |  |  | 1
| 1 | 2 | 2.8499 | 2.8555 | 2.7 | 1.71 | 2.295 | 2.856 | 2.66 | 2.76 | 2.78 | 2.78 | 2.82 | 2.75 | 2.64 | 2.8295 | 2.7593 | 2.7561 | 2.73 | 2.73 | 2.84 | 2.83 | 2.84 | 2.7006 | 3.0548 | 3.1517 |  |  | 0
| 1 | 3 | 2.8448 | 2.7959 | 2.69 | 2.01 | 2.335 | 2.855 | 2.68 | 2.65 | 2.79 | 2.68 | 2.68 | 2.66 | 2.71 | 2.8356 | 2.7285 | 2.7893 | 2.75 | 2.71 | 2.78 | 2.81 | 2.78 | 2.7529 | 2.9782 | 3.1531 |  |  | 16
| 1 | 4 | 2.9463 | 2.7867 | 2.66 | 2.03 | 2.36 | 2.867 | 2.71 | 2.8 | 2.78 | 2.79 | 3.04 | 2.71 | 2.68 | 2.5276 | 2.7742 | 2.7727 | 2.8 | 2.76 | 2.5 | 2.83 | 2.78 | 2.6962 | 2.9972 | 3.1503 |  |  | 2.7593
| 1 | 5 | 2.85 | 2.814 | 2.66 | 2.18 | 2.2967 | 2.854 | 2.72 | 2.8 | 2.78 | 2.64 | 2.88 | 2.76 | 2.73 | 2.7724 | 2.7195 | 2.7395 | 2.71 | 2.71 | 2.55 | 2.8 | 2.73 | 2.7234 | 2.9427 | 3.1502 |  |  | 36
| 1 | 6 | 2.9267 | 2.8266 | 2.74 | 1.99 | 2.28 | 2.861 | 2.73 | 2.76 | 2.78 | 2.68 | 2.84 | 2.78 | 2.68 | 2.857 | 2.7372 | 2.7644 | 2.79 | 2.74 | 2.56 | 2.84 | 2.8 | 2.7168 | 2.9453 | 3.1497
| Mean |  | 2.8849 | 2.815 | 2.6883 | 1.9233 | 2.3182 | 2.8598 | 2.6967 | 2.77 | 2.7833 | 2.7233 | 2.8133 | 2.7283 | 2.6867 | 2.7603 | 2.7405 | 2.7672 | 2.745 | 2.7283 | 2.6083 | 2.8167 | 2.7767 | 2.7174 | 2.9994 | 3.1514
| Median |  | 2.871 | 2.8125 | 2.685 | 2 | 2.3158 | 2.8585 | 2.695 | 2.78 | 2.78 | 2.725 | 2.83 | 2.73 | 2.68 | 2.8009 | 2.7328 | 2.7685 | 2.74 | 2.725 | 2.555 | 2.82 | 2.78 | 2.7156 | 2.9877 | 3.151
| Std Dev. |  | 0.04386 | 0.024308 | 0.029944 | 0.21295 | 0.031863 | 0.0057067 | 0.027325 | 0.067528 | 0.005164 | 0.064083 | 0.14949 | 0.043551 | 0.030768 | 0.12209 | 0.021646 | 0.01793 | 0.043704 | 0.019408 | 0.16497 | 0.019664 | 0.042269 | 0.020186 | 0.056374 | 0.0015332
| Rel.Std.Dev. |  | 1.52% | 0.8635% | 1.114% | 11.07% | 1.374% | 0.1995% | 1.013% | 2.438% | 0.1855% | 2.353% | 5.314% | 1.596% | 1.145% | 4.423% | 0.7899% | 0.648% | 1.592% | 0.7113% | 6.325% | 0.6981% | 1.522% | 0.7428% | 1.88% | 0.04865%
| PDM3 |  | 4.555% | 2.018% | -2.571% | -30.3% | -15.99% | 3.644% | -2.269% | 0.3888% | 0.872% | -1.302% | 1.959% | -1.121% | -2.631% | 0.03724% | -0.6809% | 0.2861% | -0.5173% | -1.121% | -5.47% | 2.08% | 0.6304% | -1.518% | 8.703% | 14.21%
| Z-Score (Absolute) |  | 1.53 | 0.72 | 0.74 | 9.59 | 5.02 | 1.24 | 0.65 | 0.2 | 0.35 | 0.34 | 0.7 | 0.28 | 0.76 | 0.09 | 0.14 | 0.17 | 0.09 | 0.28 | 1.67 | 0.74 | 0.28 | 0.41 | 2.85 | 4.61
| Analytical results for La in OREAS 235b (Certified Value 40.6 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 42.24 | 46.3639 | 42.5 | 44 | 35.63 | 35.3 | 41 | 42.3 | 39.2 | 42.3 | 37.6 | 41 | 36.2499 | 42 | 37.57 | 37.8 | 37.3 | 41.7 | 40.2 | 41.9 | 34.8 | 38.7003 |  |  |  |  | 1
| 1 | 2 | 41.1 | 46.8837 | 42.1 | 44.9 | 36.89 | 38.5 | 41 | 41.1 | 40.2 | 40.4 | 39.2 | 42 | 36.2865 | 41 | 38.79 | 40.1 | 36.57 | 41.9 | 40 | 40.3 | 34.21 | 38.4748 |  |  |  |  | 28
| 1 | 3 | 42.09 | 46.3639 | 40.2 | 44.8 | 34.83 | 35.6 | 40 | 39.6 | 37.5 | 39.6 | 42 | 42.9 | 36.0423 | 42 | 39.17 | 41.5 | 37 | 41.5 | 38.4 | 42.4 | 35.08 | 38.4512 |  |  |  |  | 16
| 1 | 4 | 42.12 | 46.2622 | 41.8 | 45.6 | 35.81 | 42.4 | 38 | 40.5 | 43 | 40 | 42 | 43.4 | 35.9364 | 41 | 38.3 | 39 | 37.26 | 41.2 | 39.2 | 40.1 | 34.3 | 38.0111 |  |  |  |  | 40.6019
| 1 | 5 | 41.64 | 46.4317 | 40.4 | 46 | 34.59 | 43.7 | 39 | 41.6 | 42.1 | 42.8 | 42.8 | 40 | 36.1489 | 40 | 38.54 | 38.4 | 37.45 | 40.3 | 38.2 | 40 | 34.65 | 39.2848 |  |  |  |  | 37
| 1 | 6 | 42.38 | 46.4769 | 42.3 | 44.3 | 36.41 | 43 | 41 | 39.8 | 42 | 44.1 | 42.2 | 42.3 | 36.0891 | 42 | 38.57 | 37.9 | 38.31 | 40.7 | 40.3 | 40.4 | 34.91 | 37.3473
| Mean |  | 41.9283 | 46.4637 | 41.55 | 44.9333 | 35.6933 | 39.75 | 40 | 40.8167 | 40.6667 | 41.5333 | 40.9667 | 41.9333 | 36.1255 | 41.3333 | 38.49 | 39.1167 | 37.315 | 41.2167 | 39.3833 | 40.85 | 34.6583 | 38.3783
| Median |  | 42.105 | 46.3978 | 41.95 | 44.85 | 35.72 | 40.45 | 40.5 | 40.8 | 41.1 | 41.35 | 42 | 42.15 | 36.119 | 41.5 | 38.555 | 38.7 | 37.28 | 41.35 | 39.6 | 40.35 | 34.725 | 38.463
| Std Dev. |  | 0.47608 | 0.21823 | 0.9975 | 0.75807 | 0.88611 | 3.7888 | 1.2649 | 1.0496 | 2.082 | 1.7974 | 2.0724 | 1.2517 | 0.1311 | 0.8165 | 0.53736 | 1.4414 | 0.57719 | 0.61455 | 0.9261 | 1.0291 | 0.34371 | 0.65392
| Rel.Std.Dev. |  | 1.135% | 0.4697% | 2.401% | 1.687% | 2.483% | 9.532% | 3.162% | 2.572% | 5.12% | 4.328% | 5.059% | 2.985% | 0.3629% | 1.975% | 1.396% | 3.685% | 1.547% | 1.491% | 2.352% | 2.519% | 0.9917% | 1.704%
| PDM3 |  | 3.267% | 14.44% | 2.335% | 10.67% | -12.09% | -2.098% | -1.483% | 0.5289% | 0.1594% | 2.294% | 0.8983% | 3.279% | -11.03% | 1.801% | -5.202% | -3.658% | -8.096% | 1.514% | -3.001% | 0.611% | -14.64% | -5.477%
| Z-Score (Absolute) |  | 0.67 | 3.24 | 0.46 | 2.38 | 2.86 | 0.56 | 0.42 | 0.04 | 0.04 | 0.45 | 0.13 | 0.68 | 2.62 | 0.34 | 1.28 | 0.92 | 1.94 | 0.27 | 0.77 | 0.06 | 3.45 | 1.34
| Analytical results for Li in OREAS 235b (Certified Value 55 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
1
| 1 | 1 | 55 | 57.2846 | 60 | 59 | 53.25 | 50 | 55.9 | 58.2 | 80.5 | 56.2 | 46.9 | 58.3 | 49.2 | 56.8 | 57.7562 | 56 | 45.48 | 50.8 | 58.13 | 53.5 | 53.1 | 55.6 | 59.1 | 45.735 |  |  | 1
| 1 | 2 | 53.7 | 57.5056 | 60 | 59 | 52.38 | 50.1 | 52.4 | 52.6 | 78 | 56.7 | 49.7 | 56.8 | 48 | 56.4 | 56.0862 | 56 | 48.55 | 55.6 | 58.88 | 55.4 | 52.2 | 54.1 | 59.1 | 41.6809 |  |  | 0
| 1 | 3 | 55.4 | 57.1586 | 60 | 59 | 51.8 | 50 | 54.5 | 51 | 81.9 | 55.9 | 47.5 | 55.2 | 49.6 | 57 | 56.8782 | 57 | 45.71 | 51.8 | 58.56 | 53.5 | 53.4 | 55.2 | 60 | 45.8086 |  |  | 16
| 1 | 4 | 56.1 | 56.6592 | 60 | 60 | 52.53 | 52.6 | 57.9 | 52.8 | 76.2 | 58.5 | 52.4 | 56.4 | 48.9 | 57.3 | 57.0396 | 56 | 46.93 | 52.7 | 58.46 | 54.3 | 53.4 | 55.1 | 58.8 | 44.5054 |  |  | 54.5974
| 1 | 5 | 54.4 | 57.445 | 60 | 59 | 53.7033 | 50.4 | 58 | 51.6 | 77.1 | 54.9 | 53.8 | 60.5 | 50 | 56.6 | 58.163 | 55 | 46.9 | 56.2 | 58.83 | 52.2 | 52.9 | 55 | 59.5 | 48.2844 |  |  | 38
| 1 | 6 | 56 | 56.5258 | 60 | 59 | 54.58 | 52.6 | 58.2 | 53.3 | 78.1 | 55.3 | 52.5 | 51.4 | 49.6 | 57.6 | 57.676 | 56 | 48.79 | 50.3 | 57.17 | 54.6 | 54.9 | 56.7 | 59.8 | 42.6001
| Mean |  | 55.1 | 57.0965 | 60 | 59.1667 | 53.0406 | 50.95 | 56.15 | 53.25 | 78.6333 | 56.25 | 50.4667 | 56.4333 | 49.2167 | 56.95 | 57.2665 | 56 | 47.06 | 52.9 | 58.3383 | 53.9167 | 53.3167 | 55.2833 | 59.3833 | 44.7691
| Median |  | 55.2 | 57.2216 | 60 | 59 | 52.89 | 50.25 | 56.9 | 52.7 | 78.05 | 56.05 | 51.05 | 56.6 | 49.4 | 56.9 | 57.3578 | 56 | 46.915 | 52.25 | 58.51 | 53.9 | 53.25 | 55.15 | 59.3 | 45.1202
| Std Dev. |  | 0.93381 | 0.41111 | 0 | 0.40825 | 1.0089 | 1.2865 | 2.3467 | 2.5657 | 2.15 | 1.274 | 2.8668 | 3.0664 | 0.70545 | 0.44609 | 0.7492 | 0.63246 | 1.3838 | 2.4739 | 0.63364 | 1.1053 | 0.89312 | 0.85186 | 0.46224 | 2.3951
| Rel.Std.Dev. |  | 1.695% | 0.72% | 0% | 0.69% | 1.902% | 2.525% | 4.179% | 4.818% | 2.734% | 2.265% | 5.681% | 5.434% | 1.433% | 0.7833% | 1.308% | 1.129% | 2.94% | 4.676% | 1.086% | 2.05% | 1.675% | 1.541% | 0.7784% | 5.35%
| PDM3 |  | 0.9206% | 4.577% | 9.895% | 8.369% | -2.851% | -6.681% | 2.844% | -2.468% | 44.02% | 3.027% | -7.566% | 3.363% | -9.855% | 4.309% | 4.889% | 2.569% | -13.81% | -3.109% | 6.852% | -1.247% | -2.346% | 1.256% | 8.766% | -18%
| Z-Score (Absolute) |  | 0.06 | 0.6 | NA | 1.29 | 0.74 | 1.44 | 0.29 | 0.67 | 7.74 | 0.32 | 1.6 | 0.38 | 2.01 | 0.55 | 0.66 | 0.24 | 2.73 | 0.79 | 1.01 | 0.45 | 0.65 | 0 | 1.36 | 3.49
Results from laboratory 1.05 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 235b (Certified Value 0.27 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.20; 1.08: 1.21; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.27778; 1.03: 0.32; 1.05: 0.4; 1.06: 0.3; 1.07: 0.234; 1.08: 0.28; 1.09: 0.36; 1.10: 0.3; 1.11: 0.3966; 1.12: 0.41085 |  |  | 1
| 1 | 2 | 1.01: 0.27346; 1.03: 0.32; 1.05: 0.3; 1.06: 0.3; 1.07: 0.245; 1.08: 0.29; 1.09: 0.35; 1.10: 0.2; 1.11: 0.1471; 1.12: 0.42106 |  |  | 29
| 1 | 3 | 1.01: 0.28567; 1.03: 0.32; 1.05: 0.3; 1.06: 0.3; 1.07: 0.247; 1.08: 0.29; 1.09: 0.35; 1.10: 0.3; 1.11: 0.17; 1.12: 0.44095 |  |  | 16
| 1 | 4 | 1.01: 0.27389; 1.03: 0.32; 1.05: 0.3; 1.06: 0.3; 1.07: 0.236; 1.08: 0.28; 1.09: 0.35; 1.10: 0.2; 1.11: 0.1667; 1.12: 0.41198 |  |  | 0.27208
| 1 | 5 | 1.01: 0.26477; 1.03: 0.32; 1.05: 0.4; 1.06: 0.3; 1.07: 0.244; 1.08: 0.27; 1.09: 0.35; 1.10: 0.2; 1.11: 0.1851; 1.12: 0.40013 |  |  | 39
| 1 | 6 | 1.01: 0.26073; 1.03: 0.32; 1.05: 0.3; 1.06: 0.3; 1.07: 0.242; 1.08: 0.27; 1.09: 0.33; 1.10: 0.2; 1.11: 0.1815; 1.12: 0.4046
| Mean |  | 1.01: 0.27271; 1.03: 0.32; 1.05: 0.33333; 1.06: 0.3; 1.07: 0.24133; 1.08: 0.28; 1.09: 0.34833; 1.10: 0.23333; 1.11: 0.20783; 1.12: 0.41493
| Median |  | 1.01: 0.27367; 1.03: 0.32; 1.05: 0.3; 1.06: 0.3; 1.07: 0.243; 1.08: 0.28; 1.09: 0.35; 1.10: 0.2; 1.11: 0.17575; 1.12: 0.41141
| Std Dev. |  | 1.01: 0.0089655; 1.03: 0; 1.05: 0.05164; 1.06: 0; 1.07: 0.0052026; 1.08: 0.0089443; 1.09: 0.0098319; 1.10: 0.05164; 1.11: 0.09344; 1.12: 0.014599
| Rel.Std.Dev. |  | 1.01: 3.287%; 1.03: 0%; 1.05: 15.49%; 1.06: 0%; 1.07: 2.156%; 1.08: 3.194%; 1.09: 2.823%; 1.10: 22.13%; 1.11: 44.96%; 1.12: 3.519%
| PDM3 |  | 1.01: 0.2345%; 1.03: 17.61%; 1.05: 22.51%; 1.06: 10.26%; 1.07: -11.3%; 1.08: 2.912%; 1.09: 28.03%; 1.10: -14.24%; 1.11: -23.61%; 1.12: 52.5%
| Z-Score (Absolute) |  | 1.01: 0.12; 1.03: 0.67; 1.05: NA; 1.06: NA; 1.07: 0.65; 1.08: 0; 1.09: 1.15; 1.10: NA; 1.11: 1.22; 1.12: 2.27
| Analytical results for Mg in OREAS 235b (Certified Value 1.78 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 1.7272 | 1.8261 | 1.78 | 1.72 | 1.74 | 1.874 | 1.65 | 1.84 | 1.75 | 1.74 | 1.62 | 1.71 | 1.68 | 1.7581 | 1.8095 | 1.8514 | 1.77 | 1.87 | 1.86 | 1.75 | 1.83 | 1.7869 | 2.1251 | 1.5243 |  |  | 1
| 1 | 2 | 1.6901 | 1.858 | 1.83 | 1.64 | 1.7738 | 1.847 | 1.64 | 1.79 | 1.75 | 1.73 | 1.78 | 1.73 | 1.68 | 1.7552 | 1.7854 | 1.8393 | 1.78 | 1.86 | 1.89 | 1.78 | 1.78 | 1.7681 | 2.1213 | 1.6645 |  |  | 0
| 1 | 3 | 1.6962 | 1.8321 | 1.76 | 1.61 | 1.73 | 1.844 | 1.72 | 1.83 | 1.74 | 1.69 | 1.69 | 1.7 | 1.72 | 1.7601 | 1.8133 | 1.8574 | 1.79 | 1.87 | 1.86 | 1.76 | 1.77 | 1.8067 | 2.049 | 1.64 |  |  | 16
| 1 | 4 | 1.7395 | 1.8088 | 1.73 | 1.62 | 1.74 | 1.878 | 1.84 | 1.81 | 1.72 | 1.73 | 1.86 | 1.72 | 1.71 | 1.7807 | 1.7896 | 1.8454 | 1.78 | 1.87 | 1.88 | 1.77 | 1.78 | 1.7707 | 2.0712 | 1.6503 |  |  | 1.776
| 1 | 5 | 1.6993 | 1.7835 | 1.84 | 1.59 | 1.74 | 1.841 | 1.89 | 1.82 | 1.71 | 1.67 | 1.8 | 1.76 | 1.71 | 1.7767 | 1.8101 | 1.8273 | 1.78 | 1.88 | 1.86 | 1.76 | 1.73 | 1.785 | 2.0321 | 1.6478 |  |  | 40
| 1 | 6 | 1.7486 | 1.8239 | 1.85 | 1.54 | 1.77 | 1.846 | 1.86 | 1.79 | 1.76 | 1.69 | 1.76 | 1.76 | 1.71 | 1.7973 | 1.7964 | 1.8454 | 1.78 | 1.86 | 1.88 | 1.79 | 1.78 | 1.7917 | 2.0232 | 1.6621
| Mean |  | 1.7168 | 1.8221 | 1.7983 | 1.62 | 1.749 | 1.855 | 1.7667 | 1.8133 | 1.7383 | 1.7083 | 1.7517 | 1.73 | 1.7017 | 1.7713 | 1.8007 | 1.8444 | 1.78 | 1.8683 | 1.8717 | 1.7683 | 1.7783 | 1.7849 | 2.0703 | 1.6315
| Median |  | 1.7132 | 1.825 | 1.805 | 1.615 | 1.74 | 1.8465 | 1.78 | 1.815 | 1.745 | 1.71 | 1.77 | 1.725 | 1.71 | 1.7684 | 1.803 | 1.8454 | 1.78 | 1.87 | 1.87 | 1.765 | 1.78 | 1.7859 | 2.0601 | 1.649
| Std Dev. |  | 0.024812 | 0.024792 | 0.048751 | 0.059666 | 0.018221 | 0.016444 | 0.11057 | 0.020656 | 0.019408 | 0.028577 | 0.084951 | 0.025298 | 0.017224 | 0.016446 | 0.011856 | 0.010377 | 0.0063246 | 0.0075277 | 0.013292 | 0.01472 | 0.031885 | 0.014208 | 0.044128 | 0.053308
| Rel.Std.Dev. |  | 1.445% | 1.361% | 2.711% | 3.683% | 1.042% | 0.8865% | 6.259% | 1.139% | 1.116% | 1.673% | 4.85% | 1.462% | 1.012% | 0.9285% | 0.6584% | 0.5626% | 0.3553% | 0.4029% | 0.7101% | 0.8324% | 1.793% | 0.796% | 2.131% | 3.267%
| PDM3 |  | -3.333% | 2.594% | 1.257% | -8.784% | -1.522% | 4.448% | -0.5257% | 2.102% | -2.121% | -3.81% | -1.37% | -2.59% | -4.186% | -0.262% | 1.391% | 3.849% | 0.225% | 5.199% | 5.386% | -0.4319% | 0.1312% | 0.4981% | 16.57% | -8.136%
| Z-Score (Absolute) |  | 0.94 | 0.76 | 0.38 | 2.51 | 0.42 | 1.3 | 0.13 | 0.62 | 0.59 | 1.08 | 0.38 | 0.73 | 1.18 | 0.06 | 0.42 | 1.13 | 0.08 | 1.51 | 1.57 | 0.11 | 0.06 | 0.16 | 4.78 | 2.32
| Analytical results for Mn in OREAS 235b (Certified Value 0.041 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0432 | 0.041104 | 0.0432 | 0.04 | 0.041921 | 0.0403 | 0.0377 | 0.0387 | 0.04 | 0.0402 | 0.0379 | 0.0393 | 0.04 | 0.0407 | 0.041298 | 0.0387 | 0.038 | 0.03935 | 0.0414 | 0.0394 | 0.0428 | 0.042471 | 0.0435 | 0.0376 |  |  | 1
| 1 | 2 | 0.0431 | 0.041723 | 0.043 | 0.0372 | 0.041471 | 0.0404 | 0.0379 | 0.0376 | 0.0407 | 0.0401 | 0.0408 | 0.0397 | 0.0394 | 0.0408 | 0.041006 | 0.0387 | 0.0386 | 0.03976 | 0.042 | 0.0399 | 0.0417 | 0.042009 | 0.0446 | 0.0406 |  |  | 30
| 1 | 3 | 0.0429 | 0.041153 | 0.042 | 0.0387 | 0.041513 | 0.041 | 0.04 | 0.0391 | 0.0401 | 0.0385 | 0.0389 | 0.0393 | 0.0397 | 0.0407 | 0.04129 | 0.0387 | 0.0385 | 0.03937 | 0.0412 | 0.0396 | 0.0415 | 0.042702 | 0.0422 | 0.0402 |  |  | 16
| 1 | 4 | 0.043 | 0.041117 | 0.0416 | 0.0385 | 0.041967 | 0.0416 | 0.0413 | 0.038 | 0.0404 | 0.0403 | 0.0428 | 0.0397 | 0.0402 | 0.0415 | 0.041137 | 0.0387 | 0.0387 | 0.03998 | 0.0419 | 0.0399 | 0.0416 | 0.04191 | 0.0427 | 0.0406 |  |  | 0.040554
| 1 | 5 | 0.0439 | 0.041416 | 0.0428 | 0.0375 | 0.041845 | 0.0406 | 0.0402 | 0.0382 | 0.0403 | 0.0385 | 0.0414 | 0.0409 | 0.0407 | 0.0413 | 0.041018 | 0.0387 | 0.0378 | 0.03945 | 0.0416 | 0.0396 | 0.0407 | 0.042042 | 0.0418 | 0.0402 |  |  | 41
| 1 | 6 | 0.0438 | 0.041585 | 0.043 | 0.0367 | 0.041916 | 0.0416 | 0.0409 | 0.0374 | 0.0398 | 0.039 | 0.0411 | 0.0405 | 0.0402 | 0.0417 | 0.041839 | 0.0387 | 0.0378 | 0.03944 | 0.0418 | 0.0403 | 0.0415 | 0.043054 | 0.0424 | 0.0403
| Mean |  | 0.043317 | 0.04135 | 0.0426 | 0.0381 | 0.041772 | 0.040917 | 0.039667 | 0.038167 | 0.040217 | 0.039433 | 0.040483 | 0.0399 | 0.040033 | 0.041117 | 0.041265 | 0.0387 | 0.038233 | 0.039558 | 0.04165 | 0.039783 | 0.041633 | 0.042365 | 0.042867 | 0.039917
| Median |  | 0.04315 | 0.041285 | 0.0429 | 0.038 | 0.041881 | 0.0408 | 0.0401 | 0.0381 | 0.0402 | 0.03955 | 0.04095 | 0.0397 | 0.0401 | 0.04105 | 0.041214 | 0.0387 | 0.03825 | 0.039445 | 0.0417 | 0.03975 | 0.04155 | 0.042257 | 0.04255 | 0.04025
| Std Dev. |  | 0.00042622 | 0.00026545 | 0.00064498 | 0.001205 | 0.00022092 | 0.00058109 | 0.0015214 | 0.00064704 | 0.00031885 | 0.00086178 | 0.0017815 | 0.00065727 | 0.00045019 | 0.00044008 | 0.00030844 | 7.6012e-18 | 0.00041312 | 0.0002542 | 0.00030822 | 0.00031885 | 0.00067429 | 0.00045555 | 0.0010231 | 0.0011496
| Rel.Std.Dev. |  | 0.984% | 0.642% | 1.514% | 3.163% | 0.5289% | 1.42% | 3.835% | 1.695% | 0.7928% | 2.185% | 4.401% | 1.647% | 1.125% | 1.07% | 0.7475% | 1.964e-14% | 1.081% | 0.6426% | 0.74% | 0.8015% | 1.62% | 1.075% | 2.387% | 2.88%
| PDM3 |  | 6.812% | 1.962% | 5.045% | -6.051% | 3.004% | 0.8943% | -2.188% | -5.887% | -0.8318% | -2.763% | -0.1742% | -1.613% | -1.284% | 1.387% | 1.753% | -4.572% | -5.722% | -2.455% | 2.703% | -1.9% | 2.661% | 4.465% | 5.703% | -1.572%
| Z-Score (Absolute) |  | 2.09 | 0.7 | 1.58 | 1.58 | 1 | 0.4 | 0.48 | 1.54 | 0.09 | 0.65 | 0.09 | 0.32 | 0.22 | 0.54 | 0.64 | 1.15 | 1.49 | 0.56 | 0.91 | 0.4 | 0.9 | 1.42 | 1.77 | 0.3
| Analytical results for Mo in OREAS 235b (Certified Value 0.98 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 1.1 | 1 | < 1 | < 2 | 1 | 1.07 | 1.4 | 3.1 | 0.95 | 0.87 | 1.18 | 0.88 | 0.92 | 1.0293 | 1.02 | 1.119 | 0.97 | < 1 | 0.96 | 1.06 | 0.94 | 1.5781 | < 5 | 3.5145 | 0.98239 |  | 1
| 1 | 2 | 1 | 1 | < 1 | < 2 | 1 | 1 | 1.2 | 3.3 | 0.92 | 0.85 | 0.99 | 0.89 | 0.95 | 1.0287 | 1.08 | 1.048 | 1 | < 1 | 1.03 | 1.1 | 0.96 | 0.8274 | < 5 | 2.7914 | 1.002 |  | 31
| 1 | 3 | 1 | 1 | < 1 | < 2 | 1 | 1.08 | 1.1 | 3.6 | 0.91 | 0.8 | 0.96 | 0.9 | 0.94 | 0.99859 | 1.08 | 1.001 | 1.03 | < 1 | 0.94 | 1.01 | 0.91 | 0.9019 | < 5 | 2.9027 | 1.1201 |  | 16
| 1 | 4 | 1.1 | 1 | < 1 | < 2 | 0.9 | 1.03 | 0.7 | 3.2 | 1.04 | 0.95 | 0.97 | 0.92 | 1 | 1.0301 | 1.07 | 0.971 | 0.91 | < 1 | 1.1 | 1 | 0.87 | 0.9084 | < 5 | 2.4577 | 0.99452 |  | 0.9849
| 1 | 5 | 1 | 1 | < 1 | < 2 | 1 | 1.03 | 1.1 | 3.1 | 0.93 | 0.91 | 0.98 | 0.91 | 1 | 1.0238 | 1.03 | 0.953 | 1 | < 1 | 0.92 | 1.02 | 0.9 | 0.8025 | < 5 | 2.569 | 1.0159 |  | 42
| 1 | 6 | 1.1 | 1 | < 1 | < 2 | 0.9 | 1.14 | 1 | 3.7 | 0.97 | 0.92 | 1.02 | 0.91 | 1.07 | 1.0162 | 1.06 | 1.007 | 0.93 | < 1 | 1.08 | 1.03 | 0.87 | 0.7885 | < 5 | 3.3476 | 0.94315
| Mean |  | 1.05 | 1 |  |  | 0.96667 | 1.0583 | 1.0833 | 3.3333 | 0.95333 | 0.88333 | 1.0167 | 0.90167 | 0.98 | 1.0211 | 1.0567 | 1.0165 | 0.97333 |  | 1.005 | 1.0367 | 0.90833 | 0.9678 |  | 2.9305 | 1.0097
| Median |  | 1.05 | 1 |  |  | 1 | 1.05 | 1.1 | 3.25 | 0.94 | 0.89 | 0.985 | 0.905 | 0.975 | 1.0262 | 1.065 | 1.004 | 0.985 |  | 0.995 | 1.025 | 0.905 | 0.86465 |  | 2.847 | 0.99826
| Std Dev. |  | 0.054772 | 0 |  |  | 0.05164 | 0.049565 | 0.23166 | 0.2582 | 0.04761 | 0.054283 | 0.082624 | 0.01472 | 0.054772 | 0.012193 | 0.02582 | 0.059886 | 0.045898 |  | 0.075829 | 0.037238 | 0.03656 | 0.30316 |  | 0.42175 | 0.059501
| Rel.Std.Dev. |  | 5.216% | 0% |  |  | 5.342% | 4.683% | 21.38% | 7.746% | 4.994% | 6.145% | 8.127% | 1.632% | 5.589% | 1.194% | 2.444% | 5.891% | 4.716% |  | 7.545% | 3.592% | 4.025% | 31.32% |  | 14.39% | 5.893%
| PDM3 |  | 6.609% | 1.533% |  |  | -1.852% | 7.455% | 9.994% | 238.4% | -3.205% | -10.31% | 3.225% | -8.451% | -0.4979% | 3.676% | 7.286% | 3.208% | -1.175% |  | 2.04% | 5.256% | -7.774% | -1.737% |  | 197.5% | 2.516%
| Z-Score (Absolute) |  | 0.63 | NA | 7.44 | 0.11 | 0.6 | 0.75 | 1.11 | 34.1 | 0.79 | 1.82 | 0.14 | 1.55 | 0.4 | 0.2 | 0.72 | 0.13 | 0.5 | 7.44 | 0.03 | 0.43 | 1.45 | 0.58 | 21.88 | 28.2 | 0.03
Results from laboratory 1.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 235b (Certified Value 0.754 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.7676 | 0.78985 | 0.74 | 0.77 | 0.7645 | 0.75 | 0.74 | 0.788 | 0.77 | 0.74 | 0.71 | 0.74 | 0.76 | 0.74 | 0.76554 | 0.749 | 0.73 | 0.78 | 0.75 | 0.74 | 0.742 | 0.75215 | 0.9133 |  |  |  | 1
| 1 | 2 | 0.7557 | 0.80571 | 0.75 | 0.72 | 0.759 | 0.785 | 0.75 | 0.765 | 0.77 | 0.74 | 0.77 | 0.75 | 0.75 | 0.744 | 0.72628 | 0.742 | 0.74 | 0.77 | 0.76 | 0.76 | 0.755 | 0.74411 | 0.8009 |  |  |  | 0
| 1 | 3 | 0.7438 | 0.7893 | 0.74 | 0.72 | 0.75 | 0.745 | 0.73 | 0.784 | 0.75 | 0.73 | 0.74 | 0.73 | 0.77 | 0.754 | 0.73837 | 0.749 | 0.75 | 0.78 | 0.76 | 0.75 | 0.75 | 0.75911 | 0.7898 |  |  |  | 16
| 1 | 4 | 0.7621 | 0.78912 | 0.74 | 0.74 | 0.78 | 0.76 | 0.76 | 0.779 | 0.77 | 0.74 | 0.8 | 0.73 | 0.76 | 0.753 | 0.7391 | 0.742 | 0.77 | 0.77 | 0.77 | 0.76 | 0.746 | 0.74343 | 0.8154 |  |  |  | 0.75371
| 1 | 5 | 0.7624 | 0.77301 | 0.74 | 0.72 | 0.7645 | 0.75 | 0.77 | 0.778 | 0.76 | 0.72 | 0.78 | 0.76 | 0.77 | 0.752 | 0.72805 | 0.734 | 0.74 | 0.78 | 0.76 | 0.75 | 0.73 | 0.75156 | 0.7913 |  |  |  | 43
| 1 | 6 | 0.7788 | 0.79229 | 0.76 | 0.69 | 0.781 | 0.782 | 0.76 | 0.764 | 0.79 | 0.73 | 0.77 | 0.75 | 0.76 | 0.761 | 0.71252 | 0.749 | 0.76 | 0.78 | 0.76 | 0.76 | 0.735 | 0.75656 | 0.7902
| Mean |  | 0.76173 | 0.78988 | 0.745 | 0.72667 | 0.7665 | 0.762 | 0.75167 | 0.77633 | 0.76833 | 0.73333 | 0.76167 | 0.74333 | 0.76167 | 0.75067 | 0.73497 | 0.74417 | 0.74833 | 0.77667 | 0.76 | 0.75333 | 0.743 | 0.75115 | 0.81682
| Median |  | 0.76225 | 0.78958 | 0.74 | 0.72 | 0.7645 | 0.755 | 0.755 | 0.7785 | 0.77 | 0.735 | 0.77 | 0.745 | 0.76 | 0.7525 | 0.73321 | 0.7455 | 0.745 | 0.78 | 0.76 | 0.755 | 0.744 | 0.75186 | 0.7961
| Std Dev. |  | 0.011698 | 0.010411 | 0.0083666 | 0.026583 | 0.012075 | 0.017378 | 0.01472 | 0.0098522 | 0.013292 | 0.008165 | 0.031885 | 0.012111 | 0.0075277 | 0.0075277 | 0.017836 | 0.006047 | 0.01472 | 0.005164 | 0.0063246 | 0.008165 | 0.0093381 | 0.0063703 | 0.048278
| Rel.Std.Dev. |  | 1.536% | 1.318% | 1.123% | 3.658% | 1.575% | 2.281% | 1.958% | 1.269% | 1.73% | 1.113% | 4.186% | 1.629% | 0.9883% | 1.003% | 2.427% | 0.8126% | 1.967% | 0.6649% | 0.8322% | 1.084% | 1.257% | 0.8481% | 5.911%
| PDM3 |  | 1.064% | 4.798% | -1.156% | -3.589% | 1.696% | 1.099% | -0.2717% | 3.001% | 1.94% | -2.704% | 1.055% | -1.377% | 1.055% | -0.4044% | -2.486% | -1.267% | -0.714% | 3.045% | 0.8339% | -0.05059% | -1.422% | -0.3398% | 8.372%
| Z-Score (Absolute) |  | 0.63 | 2.73 | 0.62 | 1.99 | 0.98 | 0.65 | 0.12 | 1.72 | 1.12 | 1.5 | 0.62 | 0.75 | 0.62 | 0.2 | 1.37 | 0.67 | 0.37 | 1.74 | 0.5 | 0 | 0.77 | 0.16 | 4.75
| Analytical results for Nb in OREAS 235b (Certified Value 15.8 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 17.07 | 14.2513 | 2.7 | 18 | 15 | 13 | 17 | 15.8 | 16.8 | 15.2 | 16.3 | 14.9 | 16.33 | 16.7987 | 16.9 | 14 | 17.17 | 16.4 | 14.58 | 24.7431 | 16.7776 | [5 empty] | 1
| 1 | 2 | 16.71 | 14.2844 | 0.6 | 18.3 | 15.2 | 13 | 16.3 | 15 | 16.8 | 16.4 | 15.2 | 15.3 | 17.34 | 16.3207 | 17.2 | 14.5 | 17.63 | 15.8 | 14.35 | 13.0105 | 16.9081 | [5 empty] | 32
| 1 | 3 | 16.77 | 14.2265 | 0.2 | 16.37 | 15.3 | 12 | 14.8 | 15.2 | 16.4 | 15.4 | 15.4 | 15.7 | 16.98 | 16.4435 | 16.9 | 14.5 | 17.29 | 15.4 | 14 | 12.6633 | 17.0637 | [5 empty] | 16
| 1 | 4 | 17.84 | 14.2546 | 0.5 | 17.1 | 16 | 12 | 14.7 | 14.9 | 17.4 | 18.1 | 15.6 | 15.7 | 17.9 | 16.3226 | 16.7 | 14.3 | 16.74 | 15.8 | 14.06 | 12.5242 | 16.6968 | [5 empty] | 15.8254
| 1 | 5 | 17.81 | 14.2499 | 0.5 | 17.09 | 16 | 12 | 14.2 | 15.7 | 16.2 | 17 | 16.5 | 16 | 16.15 | 16.6751 | 17 | 14.5 | 16.59 | 14.8 | 13.78 | 12.8399 | 16.9384 | [5 empty] | 44
| 1 | 6 | 18.12 | 14.2094 | 0.4 | 16.74 | 15.8 | 12 | 15.1 | 15.2 | 16.7 | 16.7 | 16.4 | 15.6 | 16.11 | 16.2585 | 17 | 14 | 16.53 | 16.1 | 13.92 | 12.5909 | 16.9924
| Mean |  | 17.3867 | 14.246 | 0.81667 | 17.2667 | 15.55 | 12.3333 | 15.35 | 15.3 | 16.7167 | 16.4667 | 15.9 | 15.5333 | 16.8017 | 16.4699 | 16.95 | 14.3 | 16.9917 | 15.7167 | 14.115 | 14.7286 | 16.8962
| Median |  | 17.44 | 14.2506 | 0.5 | 17.095 | 15.55 | 12 | 14.95 | 15.2 | 16.75 | 16.55 | 15.95 | 15.65 | 16.655 | 16.3831 | 16.95 | 14.4 | 16.955 | 15.8 | 14.03 | 12.7516 | 16.9232
| Std Dev. |  | 0.61007 | 0.025725 | 0.93256 | 0.74107 | 0.43704 | 0.5164 | 1.071 | 0.36878 | 0.41191 | 1.0727 | 0.56569 | 0.38297 | 0.72849 | 0.21889 | 0.16432 | 0.24495 | 0.43956 | 0.56006 | 0.29596 | 4.9092 | 0.13644
| Rel.Std.Dev. |  | 3.509% | 0.1806% | 114.2% | 4.292% | 2.811% | 4.187% | 6.977% | 2.41% | 2.464% | 6.514% | 3.558% | 2.465% | 4.336% | 1.329% | 0.9694% | 1.713% | 2.587% | 3.563% | 2.097% | 33.33% | 0.8075%
| PDM3 |  | 9.866% | -9.98% | -94.84% | 9.107% | -1.74% | -22.07% | -3.004% | -3.32% | 5.632% | 4.052% | 0.4713% | -1.846% | 6.169% | 4.072% | 7.106% | -9.639% | 7.37% | -0.6871% | -10.81% | -6.93% | 6.766%
| Z-Score (Absolute) |  | 1.03 | 1.02 | 9.75 | 0.95 | 0.17 | NA | 0.3 | 0.33 | 0.59 | 0.43 | 0.06 | 0.18 | 0.65 | 0.43 | 0.74 | 0.98 | 0.77 | 0.06 | 1.1 | 0.7 | 0.71
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 235b (Certified Value 35.7 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.03; 1.05: 1.05; 1.06: 1.06; 1.07: 1.16; 1.08: 1.19; 1.09: 1.20; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS |  | 1
2
| 1 | 1 | 1.01: 34.2271; 1.05: 36.6; 1.06: 36.4; 1.07: 37.1; 1.08: 37; 1.09: 33.27; 1.10: 37.15; 1.11: 36.4; 1.12: 30.25; 1.13: 34.7413 |  | 1
| 1 | 2 | 1.01: 34.0586; 1.05: 37.3; 1.06: 35.9; 1.07: 41.7; 1.08: 37.5; 1.09: 34.61; 1.10: 37.61; 1.11: 35.3; 1.12: 29.65; 1.13: 35.0908 |  | 9
| 1 | 3 | 1.01: 35.1706; 1.05: 35.5; 1.06: 35.6; 1.07: 38.8; 1.08: 36.6; 1.09: 34.33; 1.10: 36.75; 1.11: 36.7; 1.12: 30.5; 1.13: 35.3672 |  | 16
| 1 | 4 | 1.01: 34.0228; 1.05: 35.5; 1.06: 36.6; 1.07: 39.6; 1.08: 36.5; 1.09: 34.19; 1.10: 36.57; 1.11: 35.2; 1.12: 29.61; 1.13: 34.2862 |  | 35.6873
| 1 | 5 | 1.01: 34.0653; 1.05: 35.7; 1.06: 36.1; 1.07: 33.1; 1.08: 36.9; 1.09: 34.38; 1.10: 35.51; 1.11: 34.9; 1.12: 30.05; 1.13: 34.6167 |  | 45
| 1 | 6 | 1.01: 34.8937; 1.05: 36.6; 1.06: 35.1; 1.07: 34.2; 1.08: 36.8; 1.09: 34.54; 1.10: 35.19; 1.11: 35.1; 1.12: 30.42; 1.13: 34.172
| Mean |  | 1.01: 34.4063; 1.05: 36.2; 1.06: 35.95; 1.07: 37.4167; 1.08: 36.8833; 1.09: 34.22; 1.10: 36.4633; 1.11: 35.6; 1.12: 30.08; 1.13: 34.7123
| Median |  | 1.01: 34.1462; 1.05: 36.15; 1.06: 36; 1.07: 37.95; 1.08: 36.85; 1.09: 34.355; 1.10: 36.66; 1.11: 35.25; 1.12: 30.15; 1.13: 34.679
| Std Dev. |  | 1.01: 0.49759; 1.05: 0.74297; 1.06: 0.54681; 1.07: 3.2896; 1.08: 0.35449; 1.09: 0.489; 1.10: 0.93916; 1.11: 0.75366; 1.12: 0.38147; 1.13: 0.45947
| Rel.Std.Dev. |  | 1.01: 1.446%; 1.05: 2.052%; 1.06: 1.521%; 1.07: 8.792%; 1.08: 0.9611%; 1.09: 1.429%; 1.10: 2.576%; 1.11: 2.117%; 1.12: 1.268%; 1.13: 1.324%
| PDM3 |  | 1.01: -3.589%; 1.05: 1.437%; 1.06: 0.7362%; 1.07: 4.846%; 1.08: 3.352%; 1.09: -4.111%; 1.10: 2.175%; 1.11: -0.2445%; 1.12: -15.71%; 1.13: -2.732%
| Z-Score (Absolute) |  | 1.01: 0.85; 1.05: 0.26; 1.06: 0.11; 1.07: 1.02; 1.08: 0.69; 1.09: 0.97; 1.10: 0.43; 1.11: 0.11; 1.12: 3.54; 1.13: 0.66
| Analytical results for Ni in OREAS 235b (Certified Value 66 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 65.6 | 61.65 | 70 | 70 | 60.56 | 66 | 65 | 64 | 69.5 | 66.4 | 61.3 | 68 | 64.8 | 62.9 | 68.866 | 69.9 | 53.51 | 67 | 62.65 | 66.1 | 69.7 | 70.3 | 64.46 | 69 | 61.1161 | 65.5576 | 1
| 1 | 2 | 65.4 | 62.24 | 70 | 68 | 61.1703 | 66 | 64 | 64 | 69.4 | 65.8 | 65.8 | 66 | 63.4 | 62.2 | 69.505 | 69 | 66.44 | 68 | 62.32 | 65.7 | 68.9 | 70.5 | 64.46 | 69 | 61.8658 | 66.5894 | 34
| 1 | 3 | 65.9 | 60.54 | 70 | 68 | 61.3687 | 67 | 64 | 65 | 69.3 | 64.1 | 66.4 | 64.6 | 66.2 | 63 | 69.3858 | 68.7 | 62.76 | 67 | 62.35 | 65.7 | 66.4 | 69 | 64.9 | 67 | 62.1228 | 65.2867 | 16
| 1 | 4 | 66.2 | 61.7 | 70 | 69 | 61.27 | 67 | 64 | 65 | 69.5 | 67.2 | 71.4 | 67 | 65.5 | 64.4 | 69.3925 | 68.9 | 63.95 | 68 | 61.93 | 63.2 | 66.9 | 68.5 | 64.13 | 65 | 60.27 | 66.8536 | 65.8957
| 1 | 5 | 65.6 | 60.43 | 70 | 69 | 62.9333 | 67 | 67 | 65 | 69.6 | 64.7 | 70.1 | 68.8 | 65.2 | 64 | 68.8719 | 69.6 | 62.38 | 66 | 61.67 | 64 | 65.5 | 68.5 | 64.57 | 66 | 65.8713 | 66.0799 | 46
| 1 | 6 | 67.1 | 61.57 | 70 | 69 | 62.035 | 67 | 64 | 64 | 69.2 | 66.6 | 69.7 | 68.3 | 65 | 64.6 | 68.413 | 69 | 61.08 | 67 | 61.66 | 62.9 | 70.1 | 68.8 | 65.12 | 66 | 58.3743 | 65.9879
| Mean |  | 65.9667 | 61.355 | 70 | 68.8333 | 61.5562 | 66.6667 | 64.6667 | 64.5 | 69.4167 | 65.8 | 67.45 | 67.1167 | 65.0167 | 63.5167 | 69.0724 | 69.1833 | 61.6867 | 67.1667 | 62.0967 | 64.6 | 67.9167 | 69.2667 | 64.6067 | 67 | 61.6034 | 66.0592
| Median |  | 65.75 | 61.61 | 70 | 69 | 61.3193 | 67 | 64 | 64.5 | 69.45 | 66.1 | 68.05 | 67.5 | 65.1 | 63.5 | 69.1288 | 69 | 62.57 | 67 | 62.125 | 64.85 | 67.9 | 68.9 | 64.515 | 66.5 | 61.4909 | 66.0339
| Std Dev. |  | 0.62183 | 0.71489 | 0 | 0.75277 | 0.8227 | 0.5164 | 1.2111 | 0.54772 | 0.1472 | 1.1883 | 3.7238 | 1.588 | 0.93041 | 0.95586 | 0.42555 | 0.46224 | 4.3949 | 0.75277 | 0.40515 | 1.4057 | 1.902 | 0.90037 | 0.35246 | 1.6733 | 2.4898 | 0.59433
| Rel.Std.Dev. |  | 0.9426% | 1.165% | 0% | 1.094% | 1.337% | 0.7746% | 1.873% | 0.8492% | 0.212% | 1.806% | 5.521% | 2.366% | 1.431% | 1.505% | 0.6161% | 0.6681% | 7.125% | 1.121% | 0.6525% | 2.176% | 2.801% | 1.3% | 0.5455% | 2.497% | 4.042% | 0.8997%
| PDM3 |  | 0.1077% | -6.891% | 6.228% | 4.458% | -6.585% | 1.17% | -1.865% | -2.118% | 5.343% | -0.1452% | 2.359% | 1.853% | -1.334% | -3.61% | 4.821% | 4.989% | -6.387% | 1.929% | -5.765% | -1.966% | 3.067% | 5.116% | -1.956% | 1.676% | -6.514% | 0.2481%
| Z-Score (Absolute) |  | 0.02 | 1.84 | 1.57 | 1.11 | 1.76 | 0.26 | 0.53 | 0.6 | 1.34 | 0.08 | 0.57 | 0.44 | 0.39 | 0.99 | 1.21 | 1.25 | 1.71 | 0.46 | 1.55 | 0.56 | 0.75 | 1.28 | 0.56 | 0.39 | 1.74 | 0.02
| Analytical results for P in OREAS 235b (Certified Value 0.076 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0747 | 0.072039 | 0.07 | 0.077 | 0.073991 | 0.082 | 0.076 | 0.076 | 0.077 | 0.077 | 0.072 | 0.076 | 0.075 | 0.0757 | 0.07698 | 0.0794 | 0.08 | 0.08 | 0.0743 | 0.0774 | 0.075 | 0.0751 | 0.0761 | 0.0849 |  |  | 1
| 1 | 2 | 0.0722 | 0.072359 | 0.065 | 0.071 | 0.073713 | 0.08 | 0.0735 | 0.074 | 0.077 | 0.077 | 0.079 | 0.076 | 0.076 | 0.0756 | 0.077787 | 0.0785 | 0.08 | 0.08 | 0.0731 | 0.079 | 0.076 | 0.0739 | 0.0787 | 0.0795 |  |  | 0
| 1 | 3 | 0.0696 | 0.071798 | 0.065 | 0.069 | 0.073991 | 0.08 | 0.075 | 0.075 | 0.081 | 0.075 | 0.075 | 0.075 | 0.077 | 0.0757 | 0.076529 | 0.0799 | 0.07 | 0.08 | 0.0751 | 0.0781 | 0.075 | 0.0748 | 0.0739 | 0.079 |  |  | 16
| 1 | 4 | 0.0724 | 0.071703 | 0.065 | 0.069 | 0.073947 | 0.082 | 0.0773 | 0.075 | 0.077 | 0.077 | 0.079 | 0.076 | 0.076 | 0.0759 | 0.075837 | 0.079 | 0.07 | 0.08 | 0.0744 | 0.0786 | 0.076 | 0.0734 | 0.0739 | 0.0799 |  |  | 0.075924
| 1 | 5 | 0.0746 | 0.072244 | 0.07 | 0.068 | 0.074068 | 0.081 | 0.0779 | 0.075 | 0.076 | 0.073 | 0.08 | 0.08 | 0.076 | 0.0761 | 0.078244 | 0.0785 | 0.07 | 0.08 | 0.0732 | 0.0785 | 0.073 | 0.0754 | 0.0726 | 0.079 |  |  | 47
| 1 | 6 | 0.0761 | 0.072159 | 0.07 | 0.068 | 0.074227 | 0.08 | 0.0777 | 0.074 | 0.077 | 0.076 | 0.079 | 0.078 | 0.076 | 0.0775 | 0.078175 | 0.0799 | 0.08 | 0.08 | 0.0734 | 0.0789 | 0.075 | 0.0753 | 0.0728 | 0.0797
| Mean |  | 0.073267 | 0.07205 | 0.0675 | 0.070333 | 0.07399 | 0.080833 | 0.076233 | 0.074833 | 0.0775 | 0.075833 | 0.077333 | 0.076833 | 0.076 | 0.076083 | 0.077259 | 0.0792 | 0.075 | 0.08 | 0.073917 | 0.078417 | 0.075 | 0.07465 | 0.074667 | 0.080333
| Median |  | 0.0735 | 0.072099 | 0.0675 | 0.069 | 0.073991 | 0.0805 | 0.07665 | 0.075 | 0.077 | 0.0765 | 0.079 | 0.076 | 0.076 | 0.0758 | 0.077384 | 0.0792 | 0.075 | 0.08 | 0.07385 | 0.07855 | 0.075 | 0.07495 | 0.0739 | 0.0796
| Std Dev. |  | 0.0023321 | 0.00025627 | 0.0027386 | 0.0034448 | 0.00016788 | 0.00098319 | 0.0017386 | 0.00075277 | 0.0017607 | 0.0016021 | 0.0031411 | 0.0018348 | 0.00063246 | 0.00071671 | 0.00097171 | 0.00063875 | 0.0054772 | 0 | 0.00080353 | 0.00059133 | 0.0010954 | 0.0008167 | 0.0023347 | 0.0022669
| Rel.Std.Dev. |  | 3.183% | 0.3557% | 4.057% | 4.898% | 0.2269% | 1.216% | 2.281% | 1.006% | 2.272% | 2.113% | 4.062% | 2.388% | 0.8322% | 0.942% | 1.258% | 0.8065% | 7.303% | 0% | 1.087% | 0.7541% | 1.461% | 1.094% | 3.127% | 2.822%
| PDM3 |  | -3.5% | -5.102% | -11.1% | -7.363% | -2.548% | 6.466% | 0.4077% | -1.436% | 2.076% | -0.1192% | 1.856% | 1.198% | 0.1003% | 0.2101% | 1.758% | 4.315% | -1.217% | 5.369% | -2.644% | 3.283% | -1.217% | -1.678% | -1.656% | 5.808%
| Z-Score (Absolute) |  | 1.19 | 1.74 | 3.78 | 2.51 | 0.87 | 2.21 | 0.14 | 0.49 | 0.71 | 0.04 | 0.64 | 0.41 | 0.04 | 0.07 | 0.6 | 1.48 | 0.41 | 1.84 | 0.9 | 1.12 | 0.41 | 0.57 | 0.56 | 1.99
| Analytical results for Pb in OREAS 235b (Certified Value 19.1 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 20 | 19 | 16 | 16.6087 | 19 | 13.8 | 21 | 17.4 | 19.6 | 17.8 | 19.8 | 19 | 19.99 | 21.0409 | 19 | 18.44 | 19.1 | 18.03 | 20.3 | 19.9 | 20.59 | 17.9 | 15 | 9.796 | 19.3851 |  | 1
| 1 | 2 | 19.6 | 20 | 17 | 17.8487 | 19 | 11.4 | 20 | 16.5 | 19.2 | 18.1 | 18.8 | 18.8 | 19.48 | 20.5456 | 19 | 26.85 | 19.4 | 18.61 | 21 | 17.9 | 19.62 | 16.8 | 14 | 19.0983 | 19.8918 |  | 35
| 1 | 3 | 20.2 | 19 | 18 | 17.8847 | 19 | 14.5 | 19 | 17.4 | 20.7 | 18 | 19 | 19.4 | 19.27 | 21.2348 | 18.1 | 18 | 20.1 | 16.65 | 22.1 | 18.4 | 19.79 | 17.3 | 14 | 18.7003 | 19.6341 |  | 16
| 1 | 4 | 19.9 | 19 | 18 | 16.8127 | 20 | 14.2 | 19 | 16.9 | 19 | 20.7 | 21.4 | 20.3 | 19.77 | 21.3968 | 17.8 | 19.8 | 18.9 | 16.93 | 20.3 | 17.7 | 18.94 | 18.2 | 16 | 21.6353 | 19.4006 |  | 19.0867
| 1 | 5 | 19.9 | 20 | 18 | 17.6107 | 20 | 15.7 | 19 | 16.9 | 18.8 | 19.8 | 19.8 | 19.1 | 19.39 | 20.6242 | 18.6 | 17.96 | 18.3 | 18.48 | 21 | 17.7 | 19.2 | 17.3 | 14 | 16.5115 | 19.4944 |  | 48
| 1 | 6 | 20.2 | 20 | 17 | 18.348 | 19 | 13.5 | 19 | 15.3 | 18.8 | 18.6 | 19.6 | 19.5 | 19.88 | 21.3096 | 19.2 | 19 | 18.4 | 17.83 | 20.1 | 19.5 | 19 | 18.2 | 15 | 27.1072 | 19.885
| Mean |  | 19.9667 | 19.5 | 17.3333 | 17.5189 | 19.3333 | 13.85 | 19.5 | 16.7333 | 19.35 | 18.8333 | 19.7333 | 19.35 | 19.63 | 21.0253 | 18.6167 | 20.0083 | 19.0333 | 17.755 | 20.8 | 18.5167 | 19.5233 | 17.6167 | 14.6667 | 18.8081 | 19.6152
| Median |  | 19.95 | 19.5 | 17.5 | 17.7297 | 19 | 14 | 19 | 16.9 | 19.1 | 18.35 | 19.7 | 19.25 | 19.625 | 21.1379 | 18.8 | 18.72 | 19 | 17.93 | 20.65 | 18.15 | 19.41 | 17.6 | 14.5 | 18.8993 | 19.5642
| Std Dev. |  | 0.22509 | 0.54772 | 0.8165 | 0.67315 | 0.5164 | 1.4209 | 0.83666 | 0.78145 | 0.72595 | 1.1639 | 0.91797 | 0.53198 | 0.29031 | 0.36165 | 0.56006 | 3.4222 | 0.66833 | 0.80493 | 0.74297 | 0.96003 | 0.6224 | 0.57067 | 0.8165 | 5.7176 | 0.22945
| Rel.Std.Dev. |  | 1.127% | 2.809% | 4.711% | 3.842% | 2.671% | 10.26% | 4.291% | 4.67% | 3.752% | 6.18% | 4.652% | 2.749% | 1.479% | 1.72% | 3.008% | 17.1% | 3.511% | 4.534% | 3.572% | 5.185% | 3.188% | 3.239% | 5.567% | 30.4% | 1.17%
| PDM3 |  | 4.61% | 2.165% | -9.186% | -8.214% | 1.292% | -27.44% | 2.165% | -12.33% | 1.379% | -1.327% | 3.388% | 1.379% | 2.846% | 10.16% | -2.463% | 4.829% | -0.2796% | -6.977% | 8.976% | -2.987% | 2.288% | -7.702% | -23.16% | -1.46% | 2.769%
| Z-Score (Absolute) |  | 0.67 | 0.18 | 2.13 | 1.93 | 0 | 5.84 | 0.18 | 2.77 | 0.02 | 0.53 | 0.43 | 0.02 | 0.32 | 1.8 | 0.76 | 0.72 | 0.32 | 1.68 | 1.56 | 0.87 | 0.2 | 1.83 | 4.97 | 0.56 | 0.3
| Analytical results for Pd in OREAS 235b (Indicative Value 58.3 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.30 |  | ppb
|  |  | 1.01: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 182.2 |  | 1
| 1 | 2 | 1.01: 41.8 |  | 1
| 1 | 3 | 1.01: 62.2 |  | 16
| 1 | 4 | 1.01: 24.7 |  | 58.25
| 1 | 5 | 1.01: 15 |  | 11
| 1 | 6 | 1.01: 23.6
| Mean |  | 1.01: 58.25
| Median |  | 1.01: 33.25
| Std Dev. |  | 1.01: 63.0025
| Rel.Std.Dev. |  | 1.01: 108.2%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Pr in OREAS 235b (Certified Value 9.71 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.03; 1.05: 1.05; 1.06: 1.06; 1.07: 1.16; 1.08: 1.19; 1.09: 1.20; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 9.0674; 1.05: 9.7; 1.06: 9.4; 1.07: 10; 1.08: 10.5; 1.09: 9.086; 1.10: 10.13; 1.11: 10; 1.12: 7.43; 1.13: 9.2386 |  | 1
| 1 | 2 | 1.01: 8.9927; 1.05: 10.1; 1.06: 9.7; 1.07: 12.2; 1.08: 10.6; 1.09: 9.343; 1.10: 10.24; 1.11: 10; 1.12: 7.55; 1.13: 9.3173 |  | 11
| 1 | 3 | 1.01: 9.1651; 1.05: 9.6; 1.06: 9.6; 1.07: 10.5; 1.08: 10.4; 1.09: 9.235; 1.10: 10.21; 1.11: 10.2; 1.12: 7.58; 1.13: 9.4258 |  | 16
| 1 | 4 | 1.01: 9.1019; 1.05: 10; 1.06: 9.8; 1.07: 11.5; 1.08: 10.4; 1.09: 9.096; 1.10: 10.11; 1.11: 9.5; 1.12: 7.72; 1.13: 9.1675 |  | 9.7137
| 1 | 5 | 1.01: 9.0683; 1.05: 10; 1.06: 9.6; 1.07: 8.8; 1.08: 10.3; 1.09: 9.304; 1.10: 9.83; 1.11: 9.4; 1.12: 7.58; 1.13: 9.3501 |  | 49
| 1 | 6 | 1.01: 9.1501; 1.05: 9.85; 1.06: 9.4; 1.07: 9.3; 1.08: 10.4; 1.09: 9.393; 1.10: 9.94; 1.11: 9.8; 1.12: 7.73; 1.13: 9.2057
| Mean |  | 1.01: 9.0909; 1.05: 9.875; 1.06: 9.5833; 1.07: 10.3833; 1.08: 10.4333; 1.09: 9.2428; 1.10: 10.0767; 1.11: 9.8167; 1.12: 7.5983; 1.13: 9.2842
| Median |  | 1.01: 9.0851; 1.05: 9.925; 1.06: 9.6; 1.07: 10.25; 1.08: 10.4; 1.09: 9.2695; 1.10: 10.12; 1.11: 9.9; 1.12: 7.58; 1.13: 9.278
| Std Dev. |  | 1.01: 0.062988; 1.05: 0.19429; 1.06: 0.16021; 1.07: 1.2952; 1.08: 0.10328; 1.09: 0.12849; 1.10: 0.15996; 1.11: 0.31252; 1.12: 0.11268; 1.13: 0.097334
| Rel.Std.Dev. |  | 1.01: 0.6929%; 1.05: 1.968%; 1.06: 1.672%; 1.07: 12.47%; 1.08: 0.9899%; 1.09: 1.39%; 1.10: 1.587%; 1.11: 3.184%; 1.12: 1.483%; 1.13: 1.048%
| PDM3 |  | 1.01: -6.411%; 1.05: 1.661%; 1.06: -1.342%; 1.07: 6.894%; 1.08: 7.409%; 1.09: -4.847%; 1.10: 3.737%; 1.11: 1.06%; 1.12: -21.78%; 1.13: -4.422%
| Z-Score (Absolute) |  | 1.01: 0.94; 1.05: 0.27; 1.06: 0.18; 1.07: 1.06; 1.08: 1.13; 1.09: 0.71; 1.10: 0.58; 1.11: 0.18; 1.12: 3.25; 1.13: 0.64
| Analytical results for Pt in OREAS 235b (Indicative Value 7.72 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.30 |  | ppb
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 30.7 |  | 1
| 1 | 2 | 1.01: 0.8 |  | 1
| 1 | 3 | 1.01: 0.2 |  | 16
| 1 | 4 | 1.01: 2.9 |  | 7.7167
| 1 | 5 | 1.01: 2.8 |  | 12
| 1 | 6 | 1.01: 8.9
| Mean |  | 1.01: 7.7167
| Median |  | 1.01: 2.85
| Std Dev. |  | 1.01: 11.6735
| Rel.Std.Dev. |  | 1.01: 151.3%
| PDM3 |  | 1.01: -3.331e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Rb in OREAS 235b (Certified Value 159 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 167.97 | 140.1109 | 160 | 123 | 167 | 162.2 | 173 | 162.6 | 164 | 144 | 156.5 | 147 | 201.3 | 144.6072 | 165 | 139.4 | 160 | 144.9 | 158.7 | 164.8 | 178.4423 | 130.1406 |  |  |  |  | 1
| 1 | 2 | 168.25 | 142.3864 | 162 | 127 | 167.8 | 167.4 | 161 | 161.1 | 161 | 150.5 | 156.5 | 156.5 | 272.8 | 147.36 | 163 | 152 | 166 | 168.3 | 155.6 | 164.3 | 177.9612 | 117.355 |  |  |  |  | 37
| 1 | 3 | 169.13 | 141.9111 | 157 | 143 | 167.2 | 163.2 | 148 | 159.6 | 157 | 141 | 151.5 | 169.5 | 242.3 | 143.7308 | 162 | 143.5 | 174 | 163.85 | 152.5 | 163.7 | 174.7226 | 124.4793 |  |  |  |  | 16
| 1 | 4 | 171.94 | 143.4273 | 156 | 145 | 169.1 | 179.7 | 155 | 152.5 | 163.5 | 173 | 155 | 168.5 | 218.4 | 147.1537 | 162 | 146.8 | 160 | 145.35 | 152.8 | 160 | 174.4839 | 127.9116 |  |  |  |  | 158.8882
| 1 | 5 | 170.85 | 141.3556 | 161 | 149 | 168.6 | 183.3 | 155 | 153 | 162 | 166.5 | 162.5 | 172.5 | 191.5 | 145.4249 | 163 | 147.2 | 167 | 147.25 | 148.6 | 157.8 | 177.2997 | 135.2524 |  |  |  |  | 50
| 1 | 6 | 172.01 | 141.7447 | 164 | 142 | 169 | 181 | 159 | 161.4 | 154 | 165 | 165.5 | 170.5 | 195 | 145.1256 | 163 | 153.1 | 159 | 144.86 | 156.5 | 160.1 | 175.3734 | 113.6411
| Mean |  | 170.025 | 141.8227 | 160 | 138.1667 | 168.1167 | 172.8 | 158.5 | 158.3667 | 160.25 | 156.6667 | 157.9167 | 164.0833 | 220.2167 | 145.567 | 163 | 147 | 164.3333 | 152.4183 | 154.1167 | 161.7833 | 176.3805 | 124.7967
| Median |  | 169.99 | 141.8279 | 160.5 | 142.5 | 168.2 | 173.55 | 157 | 160.35 | 161.5 | 157.75 | 156.5 | 169 | 209.85 | 145.2752 | 163 | 147 | 163 | 146.3 | 154.2 | 161.9 | 176.3366 | 126.1954
| Std Dev. |  | 1.8145 | 1.1004 | 3.0332 | 10.553 | 0.91305 | 9.5789 | 8.3845 | 4.4572 | 3.9465 | 13.2426 | 5.1421 | 10.1016 | 31.857 | 1.4311 | 1.0954 | 5.1439 | 5.8195 | 10.7074 | 3.5684 | 2.8632 | 1.7294 | 8.0933
| Rel.Std.Dev. |  | 1.067% | 0.7759% | 1.896% | 7.638% | 0.5431% | 5.543% | 5.29% | 2.814% | 2.463% | 8.453% | 3.256% | 6.156% | 14.47% | 0.9831% | 0.6721% | 3.499% | 3.541% | 7.025% | 2.315% | 1.77% | 0.9805% | 6.485%
| PDM3 |  | 7.009% | -10.74% | 0.6998% | -13.04% | 5.808% | 8.756% | -0.2443% | -0.3282% | 0.8571% | -1.398% | -0.6114% | 3.27% | 38.6% | -8.384% | 2.588% | -7.482% | 3.427% | -4.072% | -3.003% | 1.822% | 11.01% | -21.46%
| Z-Score (Absolute) |  | 1.21 | 1.96 | 0.08 | 2.38 | 1 | 1.53 | 0.08 | 0.1 | 0.11 | 0.29 | 0.15 | 0.54 | 6.87 | 1.54 | 0.42 | 1.38 | 0.57 | 0.77 | 0.58 | 0.29 | 1.93 | 3.88
| Analytical results for Re in OREAS 235b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.004 | < 0.002 | 0.003 | 0.003 | < 0.002 | < 0.002 | 0.0013 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.004 | < 0.002 | 0.003 | 0.001 | < 0.002 | < 0.002 | 0.0002 | [10 empty] | 38
| 1 | 3 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.003 | < 0.002 | 0.006 | 0.003 | < 0.002 | < 0.002 | 0.0005 | [10 empty] | 16
| 1 | 4 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.004 | < 0.002 | 0.004 | 0.003 | < 0.002 | < 0.002 | 0.0008 | [10 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.003 | < 0.002 | 0.002 | 0.008 | < 0.002 | < 0.002 | 0.0013 | [10 empty] | 51
| 1 | 6 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.007 | < 0.001 | < 0.002 | < 0.002 | 0.0006
| Mean |  | 1.09: 0.0035; 1.14: 0.0036; 1.16: 0.0041667; 1.26: 0.0036; 1.30: 0.00078333
| Median |  | 1.09: 0.0035; 1.14: 0.004; 1.16: 0.0035; 1.26: 0.003; 1.30: 0.0007
| Std Dev. |  | 1.09: 0.00054772; 1.14: 0.00054772; 1.16: 0.0019408; 1.26: 0.0026077; 1.30: 0.0004446
| Rel.Std.Dev. |  | 1.09: 15.65%; 1.14: 15.21%; 1.16: 46.58%; 1.26: 72.44%; 1.30: 56.76%
PDM3
| Z-Score (Absolute) |  | 0.46 | 29.36 | 0.76 | 14.15 | 1.06 | 14.15 | 0.46 | 0.46 | 0.46 | 0.86 | 0.46 | 1.47 | 0.81 | 0.46 | 0.46 | 0.59
| Analytical results for Rh in OREAS 235b (Indicative Value 0.003 ppm) | [27 empty] | Indicative
Rh | Round | Replicate | 1.01: Lab | [10 empty] | 1
| No | No | 1.01: 1.30 | [10 empty] | ppm
|  |  | 1.01: 4A*MS | [10 empty] | 3
3
| 1 | 1 | 1.01: 0.0045 | [10 empty] | 1
| 1 | 2 | 1.01: 0.0013 | [10 empty] | 13
| 1 | 3 | 1.01: 0.0027 | [10 empty] | 16
| 1 | 4 | 1.01: 0.0021 | [10 empty] | 0.0025
| 1 | 5 | 1.01: 0.003 | [10 empty] | 9
| 1 | 6 | 1.01: 0.0014
| Mean |  | 1.01: 0.0025
| Median |  | 1.01: 0.0024
| Std Dev. |  | 1.01: 0.0011916
| Rel.Std.Dev. |  | 1.01: 47.67%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ru in OREAS 235b (Indicative Value 0.002 ppm) | [27 empty] | Indicative
Ru | Round | Replicate | 1.01: Lab | [10 empty] | 1
| No | No | 1.01: 1.30 | [10 empty] | ppm
|  |  | 1.01: 4A*MS | [10 empty] | 3
3
| 1 | 1 | 1.01: 0.0013 | [10 empty] | 1
| 1 | 2 | 1.01: 0.0061 | [10 empty] | 14
| 1 | 3 | 1.01: 0.0029 | [10 empty] | 16
| 1 | 4 | 1.01: 0.0004 | [10 empty] | 0.0023833
| 1 | 5 | 1.01: < 0.0044 | [10 empty] | 10
| 1 | 6 | 1.01: < 0.0028
| Mean |  | 1.01: 0.002675
| Median |  | 1.01: 0.0021
| Std Dev. |  | 1.01: 0.0025065
| Rel.Std.Dev. |  | 1.01: 93.7%
| PDM3 |  | 1.01: 12.24%
Z-Score (Absolute) | bar: 1.01
| Analytical results for S in OREAS 235b (Certified Value 0.141 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.14 | 0.145 | 0.14 | 0.135 | 0.151 | 0.11 | 0.145 | 0.13 | 0.14 | 0.15 | 0.14 | 0.14 | 0.1514 | 0.142 | 0.14 | 0.14 | 0.14 | 0.14 | 0.14 | 0.133 | 0.1413 | [5 empty] | 1
| 1 | 2 | 0.13 | 0.14 | 0.14 | 0.125 | 0.149 | 0.11 | 0.141 | 0.13 | 0.14 | 0.15 | 0.14 | 0.14 | 0.1452 | 0.146 | 0.15 | 0.13 | 0.14 | 0.15 | 0.14 | 0.13034 | 0.6478 | [5 empty] | 23
| 1 | 3 | 0.14 | 0.145 | 0.14 | 0.135 | 0.149 | 0.11 | 0.144 | 0.14 | 0.14 | 0.15 | 0.13 | 0.14 | 0.1442 | 0.143 | 0.15 | 0.13 | 0.15 | 0.14 | 0.14 | 0.13338 | 0.1716 | [5 empty] | 16
| 1 | 4 | 0.15 | 0.14 | 0.14 | 0.14 | 0.147 | 0.12 | 0.147 | 0.12 | 0.14 | 0.16 | 0.13 | 0.14 | 0.1452 | 0.141 | 0.16 | 0.13 | 0.14 | 0.14 | 0.13 | 0.13253 | 0.2641 | [5 empty] | 0.1409
| 1 | 5 | 0.15 | 0.135 | 0.14 | 0.14 | 0.147 | 0.13 | 0.142 | 0.12 | 0.14 | 0.16 | 0.14 | 0.14 | 0.1524 | 0.136 | 0.14 | 0.13 | 0.15 | 0.14 | 0.14 | 0.13186 | 0.1492 | [5 empty] | 52
| 1 | 6 | 0.14 | 0.14 | 0.14 | 0.15 | 0.15 | 0.12 | 0.141 | 0.13 | 0.14 | 0.15 | 0.14 | 0.14 | 0.1494 | 0.14 | 0.15 | 0.13 | 0.15 | 0.15 | 0.13 | 0.13243 | 0.1437
| Mean |  | 0.14167 | 0.14083 | 0.14 | 0.1375 | 0.14883 | 0.11667 | 0.14333 | 0.12833 | 0.14 | 0.15333 | 0.13667 | 0.14 | 0.14797 | 0.14133 | 0.14833 | 0.13167 | 0.145 | 0.14333 | 0.13667 | 0.13226 | 0.25295
| Median |  | 0.14 | 0.14 | 0.14 | 0.1375 | 0.149 | 0.115 | 0.143 | 0.13 | 0.14 | 0.15 | 0.14 | 0.14 | 0.1473 | 0.1415 | 0.15 | 0.13 | 0.145 | 0.14 | 0.14 | 0.13248 | 0.1604
| Std Dev. |  | 0.0075277 | 0.0037639 | 0 | 0.0082158 | 0.0016021 | 0.008165 | 0.0024221 | 0.0075277 | 0 | 0.005164 | 0.005164 | 0 | 0.0035495 | 0.0033267 | 0.0075277 | 0.0040825 | 0.0054772 | 0.005164 | 0.005164 | 0.0010721 | 0.1989
| Rel.Std.Dev. |  | 5.314% | 2.673% | 0% | 5.975% | 1.076% | 6.999% | 1.69% | 5.866% | 0% | 3.368% | 3.779% | 0% | 2.399% | 2.354% | 5.075% | 3.101% | 3.777% | 3.603% | 3.779% | 0.8107% | 78.63%
| PDM3 |  | 0.5458% | -0.04567% | -0.6371% | -2.411% | 5.632% | -17.2% | 1.729% | -8.917% | -0.6371% | 8.826% | -3.003% | -0.6371% | 5.017% | 0.3092% | 5.277% | -6.552% | 2.912% | 1.729% | -3.003% | -6.133% | 79.53%
| Z-Score (Absolute) |  | 0.13 | 0 | 0.13 | 0.54 | 1.29 | 3.91 | 0.4 | 2.02 | 0.13 | 2.02 | 0.67 | 0.13 | 1.15 | 0.08 | 1.21 | 1.48 | 0.67 | 0.4 | 0.67 | 1.39 | 18.14
| Analytical results for Sb in OREAS 235b (Certified Value 327 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 348.33 | 15.1 | 303.97 | 340.7 | 336.4 | 334 | 336.9 | 333 | 383 | 335 | 317 | 322.88 | 332.5188 | 264.7 | 299 | 323.3 | 308.43 | 314 | 334.9 | 338.441 | 369 | 344.8911 | 326.8246 |  |  |  | 1
| 1 | 2 | 358.02 | 12.9 | 305.5833 | 341.1 | 324.44 | 304 | 319.4 | 336 | 367 | 330 | 322 | 324.67 | 339.037 | 281.6 | 311 | 319.2 | 313.9 | 321.1 | 332.27 | 333.091 | 361 | 319.9649 | 339.4105 |  |  |  | 39
| 1 | 3 | 360.37 | 23.6 | 306.22 | 339.9 | 323.68 | 298 | 319.9 | 320 | 387 | 319 | 330 | 324.27 | 323.0437 | 269.1 | 325 | 321.5 | 303.48 | 318.5 | 329.07 | 339.618 | 358 | 326.5786 | 336.0316 |  |  |  | 16
| 1 | 4 | 356.38 | 18.9 | 307.805 | 342.8 | 339.34 | 306 | 303.6 | 336 | 336 | 325 | 326 | 324.98 | 332.9855 | 279.3 | 306 | 319 | 314.15 | 323.2 | 321.81 | 341.33 | 370 | 320.1184 | 329.866 |  |  |  | 327.4542
| 1 | 5 | 352.41 | 17.5 | 307.23 | 339.5 | 330.9 | 290 | 313.9 | 332 | 389 | 340 | 338 | 328.17 | 325.2647 | 273.5 | 317 | 324.8 | 324.11 | 323.4 | 377.82 | 332.663 | 347 | 313.2488 | 332.6158 |  |  |  | 53
| 1 | 6 | 356.65 | 21.1 | 302.49 | 343 | 338.64 | 311 | 310.2 | 324 | 370 | 332 | 327 | 338.2 | 344.5288 | 287.5 | 301 | 317.8 | 326.78 | 322.6 | 326.94 | 342.507 | 346 | 315.0782 | 342.7775
| Mean |  | 355.36 | 18.1833 | 305.5497 | 341.1667 | 332.2333 | 307.1667 | 317.3167 | 330.1667 | 372 | 330.1667 | 326.6667 | 327.195 | 332.8964 | 275.95 | 309.8333 | 320.9333 | 315.1417 | 320.4667 | 337.135 | 337.9417 | 358.5 | 323.3133 | 334.5877
| Median |  | 356.515 | 18.2 | 305.9017 | 340.9 | 333.65 | 305 | 316.65 | 332.5 | 376.5 | 331 | 326.5 | 324.825 | 332.7521 | 276.4 | 308.5 | 320.35 | 314.025 | 321.85 | 330.67 | 339.0295 | 359.5 | 320.0417 | 334.3237
| Std Dev. |  | 4.3129 | 3.9051 | 2.0108 | 1.4583 | 6.9941 | 14.9989 | 11.3577 | 6.6458 | 19.799 | 7.414 | 7.1461 | 5.666 | 8.1054 | 8.4384 | 9.9281 | 2.7362 | 8.9381 | 3.6528 | 20.4341 | 4.166 | 10.3682 | 11.5481 | 5.9811
| Rel.Std.Dev. |  | 1.214% | 21.48% | 0.6581% | 0.4274% | 2.105% | 4.883% | 3.579% | 2.013% | 5.322% | 2.246% | 2.188% | 1.732% | 2.435% | 3.058% | 3.204% | 0.8526% | 2.836% | 1.14% | 6.061% | 1.233% | 2.892% | 3.572% | 1.788%
| PDM3 |  | 8.522% | -94.45% | -6.689% | 4.188% | 1.459% | -6.196% | -3.096% | 0.8284% | 13.6% | 0.8284% | -0.2405% | -0.07915% | 1.662% | -15.73% | -5.381% | -1.991% | -3.76% | -2.134% | 2.956% | 3.203% | 9.481% | -1.265% | 2.178%
| Z-Score (Absolute) |  | 1.91 | 20.96 | 1.47 | 0.95 | 0.34 | 1.36 | 0.67 | 0.2 | 3.04 | 0.2 | 0.04 | 0 | 0.39 | 3.48 | 1.18 | 0.42 | 0.82 | 0.46 | 0.67 | 0.73 | 2.12 | 0.26 | 0.5
| Analytical results for Sc in OREAS 235b (Certified Value 14.7 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.8 | 13.5281 | 16 | 15 | 11.4527 | 15 | 13 | 15 | 14.9 | 16.1 | 14.6 | 15.4 | 14.8 | 13.5 | 14.7878 | 15.1 | 13.54 | 13.5 | 15.6 | 15.4 | 14.9 | 14.8365 | 14.5988 | 11.2974 |  |  | 1
| 1 | 2 | 14.7 | 13.361 | 17 | 15 | 10.288 | 16 | 12 | 14 | 14.7 | 16 | 15.4 | 14.8 | 14.9 | 13.3 | 14.6241 | 15 | 13.61 | 14.3 | 15 | 13.8 | 14.8 | 14.763 | 15.4559 | 11.1397 |  |  | 40
| 1 | 3 | 15 | 13.5811 | 16 | 15 | 10.2273 | 15 | 12 | 15 | 14.8 | 15.4 | 14.8 | 14.6 | 15.8 | 13.4 | 14.5451 | 14.4 | 14.12 | 15.1 | 15.6 | 14.3 | 15.4 | 14.952 | 15.4559 | 11.313 |  |  | 16
| 1 | 4 | 15.1 | 13.584 | 16 | 15 | 11.3787 | 16 | 14 | 14 | 13.8 | 16.3 | 16.4 | 15.1 | 15.5 | 13.4 | 14.5264 | 14.5 | 12.83 | 14.1 | 14.8 | 13.6 | 15.2 | 14.742 | 15.4559 | 11.3846 |  |  | 14.7174
| 1 | 5 | 15.1 | 13.8883 | 17 | 15 | 11.6607 | 15 | 13 | 14 | 14 | 15.5 | 16.5 | 15.6 | 15.6 | 13.4 | 14.8664 | 14.6 | 12.85 | 14.4 | 14.7 | 13.5 | 14.6 | 14.847 | 15.2844 | 11.3524 |  |  | 54
| 1 | 6 | 15.3 | 13.4961 | 16 | 15 | 11.012 | 15 | 13 | 14 | 14.8 | 15.6 | 15.8 | 16 | 16 | 13.9 | 14.6059 | 14.4 | 13.42 | 13.7 | 14.2 | 15.1 | 14.6 | 14.9625 | 15.6273 | 10.8049
| Mean |  | 15 | 13.5731 | 16.3333 | 15 | 11.0032 | 15.3333 | 12.8333 | 14.3333 | 14.5 | 15.8167 | 15.5833 | 15.25 | 15.4333 | 13.4833 | 14.6593 | 14.6667 | 13.395 | 14.1833 | 14.9833 | 14.2833 | 14.9167 | 14.8505 | 15.313 | 11.2153
| Median |  | 15.05 | 13.5546 | 16 | 15 | 11.1953 | 15 | 13 | 14 | 14.75 | 15.8 | 15.6 | 15.25 | 15.55 | 13.4 | 14.615 | 14.55 | 13.48 | 14.2 | 14.9 | 14.05 | 14.85 | 14.8417 | 15.4559 | 11.3052
| Std Dev. |  | 0.21909 | 0.17463 | 0.5164 | 0 | 0.61462 | 0.5164 | 0.75277 | 0.5164 | 0.47329 | 0.3656 | 0.79603 | 0.52058 | 0.48442 | 0.2137 | 0.13728 | 0.30768 | 0.49188 | 0.56716 | 0.54559 | 0.80353 | 0.32506 | 0.092177 | 0.36632 | 0.2181
| Rel.Std.Dev. |  | 1.461% | 1.287% | 3.162% | 0% | 5.586% | 3.368% | 5.866% | 3.603% | 3.264% | 2.312% | 5.108% | 3.414% | 3.139% | 1.585% | 0.9365% | 2.098% | 3.672% | 3.999% | 3.641% | 5.626% | 2.179% | 0.6207% | 2.392% | 1.945%
| PDM3 |  | 1.92% | -7.775% | 10.98% | 1.92% | -25.24% | 4.185% | -12.8% | -2.61% | -1.477% | 7.469% | 5.883% | 3.619% | 4.864% | -8.385% | -0.3953% | -0.345% | -8.986% | -3.629% | 1.807% | -2.95% | 1.354% | 0.9041% | 4.047% | -23.8%
| Z-Score (Absolute) |  | 0.29 | 1.42 | 1.88 | 0.29 | 4.49 | 0.69 | 2.3 | 0.51 | 0.31 | 1.26 | 0.99 | 0.59 | 0.81 | 1.52 | 0.12 | 0.11 | 1.63 | 0.69 | 0.27 | 0.57 | 0.19 | 0.11 | 0.66 | 4.23
| Analytical results for Se in OREAS 235b (Indicative Value 0.79 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 2 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | < 2 | < 0.5 | < 1 | 0.6 | 0.2436 | 3.524 | [7 empty] | 1
| 1 | 2 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 2 | < 1 | 2 | < 1 | < 1 | 1 | < 0.1 | < 2 | < 0.5 | < 1 | 0.6 | 0.2378 | 3.3973 | [7 empty] | 41
| 1 | 3 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 1 | < 1 | 1 | < 1 | < 1 | 1 | < 0.1 | < 2 | 0.9 | < 1 | 0.5 | 0.2414 | 2.8481 | [7 empty] | 16
| 1 | 4 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 2 | < 1 | < 1 | < 1 | < 1 | 1.1 | < 0.1 | < 2 | < 0.5 | < 1 | 0.6 | 0.2069 | 3.3973 | [7 empty] | 0.79305
| 1 | 5 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.8 | < 0.1 | < 2 | 1.2 | < 1 | 0.4 | 0.2283 | 3.3671 | [7 empty] | 11
| 1 | 6 | < 0.5 | < 5 | < 0.1 | < 5 | < 2 | < 1 | 2 | < 1 | 1 | < 1 | < 1 | 0.9 | < 0.1 | < 2 | < 0.5 | < 1 | 0.5 | 0.2117 | 3.2585
| Mean |  | 1.11: 1.8; 1.13: 1.2; 1.16: 0.95; 1.21: 1.05; 1.26: 0.53333; 1.27: 0.22828; 1.28: 3.2987
| Median |  | 1.11: 2; 1.13: 1; 1.16: 0.95; 1.21: 1.05; 1.26: 0.55; 1.27: 0.23305; 1.28: 3.3822
| Std Dev. |  | 1.11: 0.44721; 1.13: 0.44721; 1.16: 0.10488; 1.21: 0.21213; 1.26: 0.08165; 1.27: 0.015681; 1.28: 0.23645
| Rel.Std.Dev. |  | 1.11: 24.85%; 1.13: 37.27%; 1.16: 11.04%; 1.21: 20.2%; 1.26: 15.31%; 1.27: 6.869%; 1.28: 7.168%
| PDM3 |  | 1.11: 127%; 1.13: 51.32%; 1.16: 19.79%; 1.21: 32.4%; 1.26: -32.75%; 1.27: -71.21%; 1.28: 316%
| Z-Score (Absolute) |  | 0.41 | 3.09 | 0.73 | 3.09 | 0.75 | 0.03 | 1.66 | 0.03 | 0.88 | 0.03 | 0.03 | 0.67 | 0.73 | 0.75 | 0 | 0.03 | 0.03 | 0.45 | 4.33
| Analytical results for Sm in OREAS 235b (Certified Value 6.8 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.05; 1.07: 1.06; 1.08: 1.16; 1.09: 1.19; 1.10: 1.20; 1.11: 1.26; 1.12: 1.28; 1.13: 1.30; 1.14: 1.34 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.01: 6.4912; 1.06: 6.95; 1.07: 6.8; 1.08: 6; 1.09: 7.2; 1.10: 6.05; 1.11: 7.04; 1.12: 6.8; 1.13: 7.4636; 1.14: 6.4975 |  |  |  | 1
| 1 | 2 | 1.01: 6.4123; 1.06: 7.05; 1.07: 7.5; 1.08: 7.8; 1.09: 7.2; 1.10: 6.259; 1.11: 7.38; 1.12: 6.6; 1.13: 7.1252; 1.14: 6.6303 |  |  |  | 10
| 1 | 3 | 1.01: 6.5807; 1.06: 6.8; 1.07: 6.6; 1.08: 6.8; 1.09: 7.1; 1.10: 6.232; 1.11: 7.08; 1.12: 6.7; 1.13: 7.7456; 1.14: 6.6327 |  |  |  | 16
| 1 | 4 | 1.01: 6.3415; 1.06: 6.8; 1.07: 7.1; 1.08: 6.8; 1.09: 7.1; 1.10: 6.173; 1.11: 7.07; 1.12: 6.5; 1.13: 7.191; 1.14: 6.3199 |  |  |  | 6.8023
| 1 | 5 | 1.01: 6.7399; 1.06: 7.1; 1.07: 7.3; 1.08: 5.7; 1.09: 7.1; 1.10: 6.292; 1.11: 6.69; 1.12: 6.4; 1.13: 7.1346; 1.14: 6.4569 |  |  |  | 55
| 1 | 6 | 1.01: 6.6841; 1.06: 7; 1.07: 7.5; 1.08: 6; 1.09: 7.3; 1.10: 6.37; 1.11: 6.84; 1.12: 6.5; 1.13: 7.6704; 1.14: 6.4477
| Mean |  | 1.01: 6.5416; 1.06: 6.95; 1.07: 7.1333; 1.08: 6.5167; 1.09: 7.1667; 1.10: 6.2293; 1.11: 7.0167; 1.12: 6.5833; 1.13: 7.3884; 1.14: 6.4975
| Median |  | 1.01: 6.5359; 1.06: 6.975; 1.07: 7.2; 1.08: 6.4; 1.09: 7.15; 1.10: 6.2455; 1.11: 7.055; 1.12: 6.55; 1.13: 7.3273; 1.14: 6.4772
| Std Dev. |  | 1.01: 0.1552; 1.06: 0.12649; 1.07: 0.37238; 1.08: 0.77567; 1.09: 0.08165; 1.10: 0.10949; 1.11: 0.23551; 1.12: 0.1472; 1.13: 0.27764; 1.14: 0.11965
| Rel.Std.Dev. |  | 1.01: 2.373%; 1.06: 1.82%; 1.07: 5.22%; 1.08: 11.9%; 1.09: 1.139%; 1.10: 1.758%; 1.11: 3.356%; 1.12: 2.236%; 1.13: 3.758%; 1.14: 1.841%
| PDM3 |  | 1.01: -3.833%; 1.06: 2.171%; 1.07: 4.866%; 1.08: -4.2%; 1.09: 5.356%; 1.10: -8.424%; 1.11: 3.151%; 1.12: -3.22%; 1.13: 8.615%; 1.14: -4.482%
| Z-Score (Absolute) |  | 1.01: 0.58; 1.06: 0.48; 1.07: 0.95; 1.08: 0.65; 1.09: 1.04; 1.10: 1.4; 1.11: 0.65; 1.12: 0.48; 1.13: 1.62; 1.14: 0.7
| Analytical results for Sn in OREAS 235b (Certified Value 3.79 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.9 | 3.4 | 1 | < 10 | 3.9 | 3.7 | 5.1 | 4.3 | 3.5 | 3.7 | 3.9 | 3.6 | 3.9 | 4.0593 | 3.79 | 3.384 | 3.7 | < 5 | 3.7 | 3.8 | 3.5 | 5.9151 | < 5 | 3.28 | 3.8545 |  | 1
| 1 | 2 | 4 | 3.8 | 1 | < 10 | 3.8 | 3.4 | 4.6 | 4.6 | 3.5 | 3.9 | 3.8 | 3.5 | 4.1 | 3.7497 | 3.74 | 3.709 | 3.8 | < 5 | 3.8 | 3.7 | 3.5 | 3.1836 | < 5 | 3.5714 | 3.7834 |  | 26
| 1 | 3 | 4 | 3.9 | 1 | < 10 | 3.9 | 3.7 | 2.9 | 4.4 | 3.4 | 3.7 | 3.7 | 3.7 | 3.7 | 3.8565 | 3.82 | 3.533 | 4.2 | < 5 | 3.7 | 3.6 | 3.5 | 3.2419 | < 5 | 3.5913 | 3.9276 |  | 16
| 1 | 4 | 4 | 3.7 | 1 | < 10 | 3.9 | 4.2 | 3.9 | 4 | 3.6 | 4 | 3.8 | 3.7 | 4.4 | 3.85 | 3.76 | 3.598 | 3.8 | < 5 | 3.9 | 3.7 | 3.5 | 3.2701 | < 5 | 3.669 | 3.8554 |  | 3.7912
| 1 | 5 | 4.1 | 3.8 | 2 | < 10 | 3.8 | 3.8 | 3.5 | 4.2 | 3.5 | 4 | 3.9 | 3.8 | 3.8 | 4.004 | 3.72 | 3.608 | 4 | < 5 | 3.7 | 3.6 | 3.4 | 3.1981 | < 5 | 3.3183 | 3.8671 |  | 56
| 1 | 6 | 3.9 | 3.2 | 1 | < 10 | 3.9 | 4.2 | 3.7 | 4.4 | 3.5 | 4 | 4 | 3.7 | 3.9 | 3.888 | 3.7 | 3.722 | 3.7 | < 5 | 3.9 | 3.7 | 3.5 | 3.0376 | < 5 | 3.3815 | 3.8985
| Mean |  | 3.9833 | 3.6333 | 1.1667 |  | 3.8667 | 3.8333 | 3.95 | 4.3167 | 3.5 | 3.8833 | 3.85 | 3.6667 | 3.9667 | 3.9012 | 3.755 | 3.5923 | 3.8667 |  | 3.7833 | 3.6833 | 3.4833 | 3.6411 |  | 3.4686 | 3.8644
| Median |  | 4 | 3.75 | 1 |  | 3.9 | 3.75 | 3.8 | 4.35 | 3.5 | 3.95 | 3.85 | 3.7 | 3.9 | 3.8722 | 3.75 | 3.603 | 3.8 |  | 3.75 | 3.7 | 3.5 | 3.22 |  | 3.4764 | 3.8612
| Std Dev. |  | 0.075277 | 0.27325 | 0.40825 |  | 0.05164 | 0.31411 | 0.7893 | 0.20412 | 0.063246 | 0.1472 | 0.10488 | 0.10328 | 0.25033 | 0.11253 | 0.044609 | 0.1246 | 0.19664 |  | 0.098319 | 0.075277 | 0.040825 | 1.1169 |  | 0.16219 | 0.048807
| Rel.Std.Dev. |  | 1.89% | 7.521% | 34.99% |  | 1.336% | 8.194% | 19.98% | 4.729% | 1.807% | 3.79% | 2.724% | 2.817% | 6.311% | 2.884% | 1.188% | 3.468% | 5.085% |  | 2.599% | 2.044% | 1.172% | 30.68% |  | 4.676% | 1.263%
| PDM3 |  | 5.069% | -4.163% | -69.23% |  | 1.992% | 1.112% | 4.19% | 13.86% | -7.68% | 2.431% | 1.552% | -3.284% | 4.629% | 2.904% | -0.9539% | -5.245% | 1.992% |  | -0.2065% | -2.844% | -8.12% | -3.959% |  | -8.509% | 1.932%
| Z-Score (Absolute) |  | 0.79 | 0.79 | NA | 5.36 | 0.26 | 0.11 | 0.64 | 2.29 | 1.39 | 0.34 | 0.19 | 0.64 | 0.71 | 0.42 | 0.24 | 0.97 | 0.26 | 5.88 | 0.11 | 0.56 | 1.46 | 0.75 | 5.88 | 1.53 | 0.25
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 235b (Certified Value 149 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 152.81 | 153.1642 | 148 | 148 | 138.7427 | 154.5 | 140 | 169 | 147 | 151.5 | 140.5 | 150.5 | 146.5 | 147 | 148.1196 | 151 | 131.4 | 142 | 147.4 | 152.9 | 151 | 147 | 170.8 | 161 | 124.27 | 142.5807 | 1
| 1 | 2 | 153.3 | 153.7118 | 152 | 135 | 140.472 | 153.9 | 142 | 153 | 147 | 149.5 | 151.5 | 148.5 | 146 | 148 | 149.1589 | 151 | 143 | 143 | 147.3 | 151.7 | 155.3 | 147 | 170.3 | 166 | 151.1705 | 141.7668 | 27
| 1 | 3 | 151.84 | 153.5562 | 147 | 133 | 141.0882 | 153.9 | 145 | 147 | 155 | 145 | 145 | 148 | 150 | 148 | 143.3663 | 151 | 134.5 | 145 | 146.1 | 150.3 | 152.5 | 147 | 174.1 | 157 | 150.5144 | 142.6228 | 16
| 1 | 4 | 154.31 | 153.2622 | 145 | 132 | 143.0781 | 153.9 | 158 | 151 | 148 | 150.5 | 160.5 | 148 | 149.5 | 153 | 147.3791 | 150 | 137.5 | 148 | 148.3 | 148 | 153.3 | 144 | 168.9 | 156 | 151.9276 | 144.3817 | 149.4852
| 1 | 5 | 153.54 | 152.3606 | 148 | 137 | 140.2895 | 154.4 | 162 | 149 | 146 | 147 | 156.5 | 154.5 | 150 | 153 | 142.8911 | 150 | 138.4 | 142 | 146.9 | 146.7 | 152.5 | 143 | 171.5 | 154 | 149.8583 | 147.5515 | 57
| 1 | 6 | 155.48 | 156.2904 | 152 | 131 | 145.92 | 154.6 | 162 | 154 | 149 | 150 | 154 | 154 | 149 | 156 | 145.9032 | 151 | 143.5 | 148 | 145.5 | 146.5 | 153.5 | 145 | 173 | 154 | 150.363 | 139.9205
| Mean |  | 153.5467 | 153.7242 | 148.6667 | 136 | 141.5984 | 154.2 | 151.5 | 153.8333 | 148.6667 | 148.9167 | 151.3333 | 150.5833 | 148.5 | 150.8333 | 146.1364 | 150.6667 | 138.05 | 144.6667 | 146.9167 | 149.35 | 153.0167 | 145.5 | 171.4333 | 158 | 146.3506 | 143.1373
| Median |  | 153.42 | 153.4092 | 148 | 134 | 140.7801 | 154.15 | 151.5 | 152 | 147.5 | 149.75 | 152.75 | 149.5 | 149.25 | 150.5 | 146.6412 | 151 | 137.95 | 144 | 147.1 | 149.15 | 152.9 | 146 | 171.15 | 156.5 | 150.4387 | 142.6018
| Std Dev. |  | 1.2519 | 1.3417 | 2.8048 | 6.261 | 2.5402 | 0.33466 | 10.2713 | 7.8592 | 3.266 | 2.4376 | 7.4207 | 2.9903 | 1.7889 | 3.656 | 2.5641 | 0.5164 | 4.7213 | 2.8048 | 0.99683 | 2.6831 | 1.4233 | 1.7607 | 1.8801 | 4.6904 | 10.8409 | 2.6013
| Rel.Std.Dev. |  | 0.8153% | 0.8728% | 1.887% | 4.604% | 1.794% | 0.217% | 6.78% | 5.109% | 2.197% | 1.637% | 4.904% | 1.986% | 1.205% | 2.424% | 1.755% | 0.3427% | 3.42% | 1.939% | 0.6785% | 1.797% | 0.9301% | 1.21% | 1.097% | 2.969% | 7.407% | 1.817%
| PDM3 |  | 2.717% | 2.836% | -0.5476% | -9.021% | -5.276% | 3.154% | 1.348% | 2.909% | -0.5476% | -0.3803% | 1.236% | 0.7346% | -0.6591% | 0.9018% | -2.24% | 0.7904% | -7.65% | -3.223% | -1.718% | -0.09045% | 2.362% | -2.666% | 14.68% | 5.696% | -2.097% | -4.246%
| Z-Score (Absolute) |  | 0.9 | 0.93 | 0.09 | 2.67 | 1.53 | 1.03 | 0.48 | 0.96 | 0.09 | 0.04 | 0.45 | 0.29 | 0.13 | 0.35 | 0.61 | 0.31 | 2.25 | 0.91 | 0.45 | 0.04 | 0.79 | 0.74 | 4.54 | 1.8 | 0.57 | 1.22
| Analytical results for Ta in OREAS 235b (Certified Value 1.18 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 1.25 | 1.137 | 0.2 | 5.21 | 1.11 | 1.13 | 1.2 | 1 | 1.15 | 1.1 | 1.17 | 1.14 | 1.2 | 1.2851 | 1.1 | 1.11 | 1.32 | 1.24 | 1 | 0.8073 | 1.3175 | [5 empty] | 1
| 1 | 2 | 1.16 | 1.1232 | < 0.1 | < 5 | 1.1 | 1.02 | 1.19 | 1 | 1.17 | 1.14 | 1.12 | 1.12 | 1.3 | 1.3285 | 1.1 | 1.13 | 1.35 | 1.2 | 1 | 0.8117 | 1.3298 | [5 empty] | 28
| 1 | 3 | 1.2 | 1.1109 | < 0.1 | < 5 | 1.13 | 0.95 | 1.11 | 1 | 1.16 | 1.12 | 1.12 | 1.18 | 1.2 | 1.3236 | 1.09 | 1.11 | 1.31 | 1.14 | 1 | 0.8269 | 1.3212 | [5 empty] | 16
| 1 | 4 | 1.29 | 1.1258 | < 0.1 | 5.0467 | 1.16 | 1.03 | 1.05 | 0.9 | 1.2 | 1.26 | 1.12 | 1.13 | 1.2 | 1.4324 | 1.09 | 1.1 | 1.29 | 1.18 | 1 | 0.7835 | 1.304 | [5 empty] | 1.1775
| 1 | 5 | 1.26 | 1.1435 | < 0.1 | 5.5033 | 1.15 | 1.19 | 1.04 | 0.9 | 1.14 | 1.24 | 1.17 | 1.16 | 1.1 | 1.352 | 1.1 | 1.08 | 1.31 | 1.11 | 1 | 0.7378 | 1.3024 | [5 empty] | 58
| 1 | 6 | 1.29 | 1.1219 | < 0.1 | < 5 | 1.16 | 1.19 | 1.12 | 1 | 1.16 | 1.17 | 1.19 | 1.16 | 1.2 | 1.2617 | 1.1 | 1.06 | 1.26 | 1.19 | 0.9 | 0.7856 | 1.3109
| Mean |  | 1.2417 | 1.1271 | 0.2 | 5.2533 | 1.135 | 1.085 | 1.1183 | 0.96667 | 1.1633 | 1.1717 | 1.1483 | 1.1483 | 1.2 | 1.3306 | 1.0967 | 1.0983 | 1.3067 | 1.1767 | 0.98333 | 0.79213 | 1.3143
| Median |  | 1.255 | 1.1245 | 0.2 | 5.21 | 1.14 | 1.08 | 1.115 | 1 | 1.16 | 1.155 | 1.145 | 1.15 | 1.2 | 1.326 | 1.1 | 1.105 | 1.31 | 1.185 | 1 | 0.79645 | 1.3142
| Std Dev. |  | 0.051929 | 0.011588 |  | 0.2314 | 0.025884 | 0.099549 | 0.067355 | 0.05164 | 0.020656 | 0.065243 | 0.031885 | 0.022286 | 0.063246 | 0.05948 | 0.005164 | 0.024833 | 0.030111 | 0.045898 | 0.040825 | 0.031276 | 0.010577
| Rel.Std.Dev. |  | 4.182% | 1.028% |  | 4.405% | 2.281% | 9.175% | 6.023% | 5.342% | 1.776% | 5.568% | 2.777% | 1.941% | 5.27% | 4.47% | 0.4709% | 2.261% | 2.304% | 3.901% | 4.152% | 3.948% | 0.8048%
| PDM3 |  | 5.453% | -4.28% | -83.01% | 346.2% | -3.606% | -7.853% | -5.022% | -17.9% | -1.2% | -0.4921% | -2.474% | -2.474% | 1.914% | 13% | -6.862% | -6.72% | 10.97% | -0.06748% | -16.49% | -32.73% | 11.62%
| Z-Score (Absolute) |  | 1.26 | 0.29 | NA | 36.79 | 0.18 | 0.85 | 0.4 | NA | 0.2 | 0.31 | 0 | 0 | NA | 2.46 | 0.7 | 0.67 | 2.14 | 0.38 | NA | 4.8 | 2.24
Results from laboratories 1.06, 1.11, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 235b (Certified Value 0.69 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.21; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*OES; 1.13: 4A*MS | 2
3
| 1 | 1 | 1.01: 0.6763; 1.03: 0.7; 1.05: 0.7; 1.06: 0.6; 1.07: 0.71; 1.08: 0.633; 1.09: 0.62; 1.10: 0.71; 1.11: 0.6; 1.12: 2.13; 1.13: 5.5409 | 1
| 1 | 2 | 1.01: 0.718; 1.03: 0.74; 1.05: 0.7; 1.06: 0.8; 1.07: 0.7; 1.08: 0.649; 1.09: 0.61; 1.10: 0.75; 1.11: 0.6; 1.12: 1.97; 1.13: 5.62 | 11
| 1 | 3 | 1.01: 0.69245; 1.03: 0.7; 1.05: 0.8; 1.06: 0.8; 1.07: 0.7; 1.08: 0.662; 1.09: 0.64; 1.10: 0.72; 1.11: 0.6; 1.12: 2.52; 1.13: 5.7113 | 16
| 1 | 4 | 1.01: 0.69935; 1.03: 0.7; 1.05: 0.7; 1.06: 0.7; 1.07: 0.69; 1.08: 0.69; 1.09: 0.63; 1.10: 0.71; 1.11: 0.6; 1.12: 2.35; 1.13: 5.4466 | 0.68857
| 1 | 5 | 1.01: 0.72146; 1.03: 0.7; 1.05: 0.7; 1.06: 0.6; 1.07: 0.72; 1.08: 0.68; 1.09: 0.63; 1.10: 0.72; 1.11: 0.6; 1.12: 2.38; 1.13: 5.4319 | 59
| 1 | 6 | 1.01: 0.68685; 1.03: 0.72; 1.05: 0.7; 1.06: 0.6; 1.07: 0.69; 1.08: 0.669; 1.09: 0.61; 1.10: 0.69; 1.11: 0.6; 1.12: 1.87; 1.13: 5.4902
| Mean |  | 1.01: 0.69907; 1.03: 0.71; 1.05: 0.71667; 1.06: 0.68333; 1.07: 0.70167; 1.08: 0.66383; 1.09: 0.62333; 1.10: 0.71667; 1.11: 0.6; 1.12: 2.2033; 1.13: 5.5401
| Median |  | 1.01: 0.6959; 1.03: 0.7; 1.05: 0.7; 1.06: 0.65; 1.07: 0.7; 1.08: 0.6655; 1.09: 0.625; 1.10: 0.715; 1.11: 0.6; 1.12: 2.24; 1.13: 5.5156
| Std Dev. |  | 1.01: 0.017728; 1.03: 0.016733; 1.05: 0.040825; 1.06: 0.098319; 1.07: 0.01169; 1.08: 0.020721; 1.09: 0.012111; 1.10: 0.019664; 1.11: 0; 1.12: 0.25453; 1.13: 0.1083
| Rel.Std.Dev. |  | 1.01: 2.536%; 1.03: 2.357%; 1.05: 5.696%; 1.06: 14.39%; 1.07: 1.666%; 1.08: 3.121%; 1.09: 1.943%; 1.10: 2.744%; 1.11: 0%; 1.12: 11.55%; 1.13: 1.955%
| PDM3 |  | 1.01: 1.524%; 1.03: 3.112%; 1.05: 4.08%; 1.06: -0.7606%; 1.07: 1.902%; 1.08: -3.593%; 1.09: -9.474%; 1.10: 4.08%; 1.11: -12.86%; 1.12: 220%; 1.13: 704.6%
| Z-Score (Absolute) |  | 1.01: 0.1; 1.03: 0.31; 1.05: 0.55; 1.06: 0.67; 1.07: 0; 1.08: 1.39; 1.09: 2.88; 1.10: 0.55; 1.11: 3.74; 1.12: 55.23; 1.13: 177.96
| Analytical results for Te in OREAS 235b (Indicative Value 0.055 ppm) | [27 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | < 0.1 | 10.502 | < 0.1 | 0.06 | 0.4 | 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.08 | < 0.05 | 0.02695 | [8 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.1 | 9.3867 | < 0.1 | 0.07 | 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.05 | 0.07 | 0.06 | 0.02992 | [8 empty] | 6
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 10.587 | < 0.1 | 0.08 | < 0.1 | 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.04 | 0.07 | < 0.05 | 0.03995 | [8 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 10.349 | < 0.1 | 0.08 | < 0.1 | 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.06 | < 0.05 | 0.03003 | [8 empty] | 0.055267
| 1 | 5 | < 0.2 | < 0.2 | < 0.1 | 10.896 | < 0.1 | 0.1 | 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.04 | 0.06 | < 0.05 | 0.02596 | [8 empty] | 12
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 9.82 | < 0.1 | 0.1 | 0.2 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.07 | 0.05 | 0.04499
| Mean |  | 1.07: 10.2568; 1.09: 0.081667; 1.10: 0.2; 1.11: 0.1; 1.15: 0.05; 1.26: 0.043333; 1.27: 0.068333; 1.28: 0.055; 1.30: 0.032967
| Median |  | 1.07: 10.4255; 1.09: 0.08; 1.10: 0.15; 1.11: 0.1; 1.15: 0.05; 1.26: 0.04; 1.27: 0.07; 1.28: 0.055; 1.30: 0.029975
| Std Dev. |  | 1.07: 0.55392; 1.09: 0.016021; 1.10: 0.14142; 1.11: 1.6997e-17; 1.15: 7.6012e-18; 1.26: 0.0057735; 1.27: 0.0075277; 1.28: 0.0070711; 1.30: 0.007701
| Rel.Std.Dev. |  | 1.07: 5.401%; 1.09: 19.62%; 1.10: 70.71%; 1.11: 1.7e-14%; 1.15: 1.52e-14%; 1.26: 13.32%; 1.27: 11.02%; 1.28: 12.86%; 1.30: 23.36%
| PDM3 |  | 1.07: 18459%; 1.09: 47.77%; 1.10: 261.9%; 1.11: 80.94%; 1.15: -9.53%; 1.26: -21.59%; 1.27: 23.64%; 1.28: -0.4825%; 1.30: -40.35%
| Z-Score (Absolute) |  | 1.35 | 1.35 | 0 | 275.3 | 0 | 0.85 | 2.7 | 0.67 | 0.67 | 0.67 | 0.67 | 0 | 0.67 | 0.67 | 0.56 | 0.49 | 0.4 | 0.46
| Analytical results for Th in OREAS 235b (Certified Value 14.8 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 15.36 | 15.286 | 17.3 | 15.7 | 15.28 | 13 | 17.2 | 15.3 | 15.3 | 13.6 | 14.7 | 14.15 | 17.3 | 13.865 | 14.1 | 13.5 | 13.8 | 14.6 | 16 | 15 | 14.6176 | 13.4708 | 14.3325 |  |  |  | 1
| 1 | 2 | 15.13 | 15.4201 | 17.9 | 16.4 | 15.07 | 13 | 15.5 | 15 | 15.1 | 13.85 | 14.4 | 14.65 | 21.3 | 14.0065 | 14.1 | 14.12 | 14.6 | 14.9 | 15.9 | 14.3 | 14.4661 | 15.126 | 14.2432 |  |  |  | 31
| 1 | 3 | 15.12 | 15.4204 | 17.4 | 16.5 | 15.19 | 12 | 15 | 15.2 | 14.85 | 13.5 | 14.15 | 15.65 | 19.4 | 14.0943 | 14 | 14.43 | 15.2 | 14.6 | 15.3 | 14.9 | 14.2241 | 14.9684 | 14.2772 |  |  |  | 16
| 1 | 4 | 15.26 | 15.4923 | 17.9 | 16.6 | 15.25 | 16 | 14.9 | 14.3 | 14.85 | 15.5 | 14.65 | 15.5 | 19.5 | 13.9643 | 14.1 | 13.56 | 14.7 | 14.6 | 15.5 | 14.5 | 14.1984 | 15.2245 | 14.4523 |  |  |  | 14.8145
| 1 | 5 | 14.98 | 15.1818 | 17.5 | 16.5 | 15.14 | 15 | 15 | 14.2 | 15.2 | 15.35 | 14.55 | 15.65 | 15.5 | 13.5767 | 14 | 14.15 | 14.2 | 14.4 | 15.2 | 14.2 | 14.4882 | 14.8501 | 14.4504 |  |  |  | 60
| 1 | 6 | 15.32 | 15.279 | 18.2 | 16.2 | 15.05 | 15 | 15.4 | 14.9 | 14 | 14.2 | 14.65 | 15.75 | 16 | 14.2606 | 13.9 | 14.28 | 13.9 | 14.1 | 15.8 | 14 | 14.3013 | 15.2442 | 14.1579
| Mean |  | 15.195 | 15.3466 | 17.7 | 16.3167 | 15.1633 | 14 | 15.5 | 14.8167 | 14.8833 | 14.3333 | 14.5167 | 15.225 | 18.1667 | 13.9612 | 14.0333 | 14.0067 | 14.4 | 14.5333 | 15.6167 | 14.4833 | 14.3826 | 14.814 | 14.3189
| Median |  | 15.195 | 15.353 | 17.7 | 16.45 | 15.165 | 14 | 15.2 | 14.95 | 14.975 | 14.025 | 14.6 | 15.575 | 18.35 | 13.9854 | 14.05 | 14.135 | 14.4 | 14.6 | 15.65 | 14.4 | 14.3837 | 15.0472 | 14.3049
| Std Dev. |  | 0.14363 | 0.11618 | 0.35214 | 0.33116 | 0.093737 | 1.5492 | 0.86718 | 0.46224 | 0.46975 | 0.88072 | 0.20897 | 0.66314 | 2.266 | 0.23086 | 0.08165 | 0.38563 | 0.53292 | 0.26583 | 0.33116 | 0.39707 | 0.16674 | 0.67532 | 0.11717
| Rel.Std.Dev. |  | 0.9453% | 0.757% | 1.989% | 2.03% | 0.6182% | 11.07% | 5.595% | 3.12% | 3.156% | 6.145% | 1.439% | 4.356% | 12.47% | 1.654% | 0.5818% | 2.753% | 3.701% | 1.829% | 2.121% | 2.742% | 1.159% | 4.559% | 0.8183%
| PDM3 |  | 2.568% | 3.592% | 19.48% | 10.14% | 2.355% | -5.498% | 4.627% | 0.0145% | 0.4645% | -3.248% | -2.011% | 2.771% | 22.63% | -5.76% | -5.273% | -5.453% | -2.798% | -1.898% | 5.415% | -2.236% | -2.915% | -0.003497% | -3.345%
| Z-Score (Absolute) |  | 0.56 | 0.78 | 4.25 | 2.21 | 0.51 | NA | 1.01 | 0 | 0.1 | 0.71 | 0.44 | 0.6 | 4.94 | 1.26 | 1.15 | 1.19 | 0.61 | 0.42 | 1.18 | 0.49 | 0.64 | 0 | 0.73
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 235b (Certified Value 0.489 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.49 | 0.47311 | 0.49 | 0.458 | 0.37667 | 0.544 | 0.48 | 0.483 | 0.501 | 0.494 | 0.46 | 0.481 | 0.484 | 0.4886 | 0.4965 | 0.5503 | 0.48 | 0.49 | 0.48 | 0.53 | 0.493 | 0.41064 | 0.5333 | 0.4611 | 0.50808 |  | 1
| 1 | 2 | 0.4789 | 0.47477 | 0.49 | 0.384 | 0.38 | 0.535 | 0.49 | 0.485 | 0.504 | 0.496 | 0.504 | 0.476 | 0.479 | 0.4906 | 0.50611 | 0.5485 | 0.49 | 0.49 | 0.49 | 0.53 | 0.487 | 0.40898 | 0.5296 | 0.4642 | 0.51015 |  | 32
| 1 | 3 | 0.4871 | 0.47424 | 0.475 | 0.229 | 0.37 | 0.539 | 0.48 | 0.479 | 0.514 | 0.474 | 0.473 | 0.474 | 0.489 | 0.4978 | 0.50082 | 0.5539 | 0.49 | 0.49 | 0.48 | 0.53 | 0.48 | 0.41578 | 0.5127 | 0.4671 | 0.48985 |  | 16
| 1 | 4 | 0.504 | 0.47364 | 0.48 | 0.329 | 0.38333 | 0.559 | 0.48 | 0.475 | 0.503 | 0.49 | 0.528 | 0.475 | 0.483 | 0.5008 | 0.49486 | 0.5491 | 0.49 | 0.48 | 0.49 | 0.53 | 0.484 | 0.40483 | 0.5143 | 0.4613 | 0.49442 |  | 0.48921
| 1 | 5 | 0.501 | 0.47637 | 0.475 | 0.296 | 0.38 | 0.564 | 0.49 | 0.47 | 0.51 | 0.475 | 0.515 | 0.489 | 0.489 | 0.5029 | 0.50308 | 0.5437 | 0.48 | 0.5 | 0.49 | 0.51 | 0.473 | 0.40963 | 0.5064 | 0.4313 | 0.50611 |  | 61
| 1 | 6 | 0.5131 | 0.47985 | 0.495 | 0.267 | 0.395 | 0.549 | 0.49 | 0.464 | 0.518 | 0.48 | 0.495 | 0.492 | 0.488 | 0.5069 | 0.50147 | 0.5497 | 0.49 | 0.49 | 0.48 | 0.53 | 0.478 | 0.41703 | 0.5042 | 0.4745 | 0.50122
| Mean |  | 0.49568 | 0.47533 | 0.48417 | 0.32717 | 0.38083 | 0.54833 | 0.485 | 0.476 | 0.50833 | 0.48483 | 0.49583 | 0.48117 | 0.48533 | 0.49793 | 0.50047 | 0.5492 | 0.48667 | 0.49 | 0.485 | 0.52667 | 0.4825 | 0.41115 | 0.51675 | 0.45992 | 0.50164
| Median |  | 0.4955 | 0.47451 | 0.485 | 0.3125 | 0.38 | 0.5465 | 0.485 | 0.477 | 0.507 | 0.485 | 0.4995 | 0.4785 | 0.486 | 0.4993 | 0.50114 | 0.5494 | 0.49 | 0.49 | 0.485 | 0.53 | 0.482 | 0.41014 | 0.5135 | 0.46275 | 0.50366
| Std Dev. |  | 0.012557 | 0.0024819 | 0.008612 | 0.083161 | 0.0082832 | 0.011343 | 0.0054772 | 0.008 | 0.0067725 | 0.0097245 | 0.025592 | 0.007679 | 0.0040332 | 0.0071301 | 0.0041699 | 0.0032955 | 0.005164 | 0.0063246 | 0.0054772 | 0.008165 | 0.007064 | 0.0045432 | 0.01205 | 0.014865 | 0.0080636
| Rel.Std.Dev. |  | 2.533% | 0.5221% | 1.779% | 25.42% | 2.175% | 2.069% | 1.129% | 1.681% | 1.332% | 2.006% | 5.161% | 1.596% | 0.831% | 1.432% | 0.8332% | 0.6% | 1.061% | 1.291% | 1.129% | 1.55% | 1.464% | 1.105% | 2.332% | 3.232% | 1.607%
| PDM3 |  | 1.323% | -2.837% | -1.031% | -33.12% | -22.15% | 12.09% | -0.8604% | -2.7% | 3.909% | -0.8944% | 1.354% | -1.644% | -0.7922% | 1.783% | 2.302% | 12.26% | -0.5197% | 0.1617% | -0.8604% | 7.657% | -1.371% | -15.96% | 5.63% | -5.988% | 2.541%
| Z-Score (Absolute) |  | 0.66 | 0.64 | 0.07 | 10.16 | 6.71 | 4.05 | 0.02 | 0.6 | 1.48 | 0.03 | 0.67 | 0.27 | 0 | 0.81 | 0.97 | 4.1 | 0.09 | 0.3 | 0.02 | 2.65 | 0.18 | 4.76 | 2.02 | 1.63 | 1.05
| Analytical results for Tl in OREAS 235b (Certified Value 0.86 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.91 | 0.81978 | 0.8 | 0.88 | 0.86 | 0.67 | 0.88 | 0.9 | 0.89 | 0.83 | 0.86 | 0.84 | 0.88 | 0.88115 | 0.787 | 0.82 | 0.89 | 0.9 | 0.86 | 1.0381 | 0.8535 | 0.89903 |  |  |  |  | 1
| 1 | 2 | 0.91 | 0.82294 | 0.8 | 0.88 | 0.88 | 0.78 | 0.84 | 1 | 0.92 | 0.84 | 0.82 | 0.83 | 0.86 | 0.88758 | 0.837 | 0.85 | 0.89 | 0.8 | 0.83 | 1.0284 | 0.9041 | 0.90875 |  |  |  |  | 33
| 1 | 3 | 0.89 | 0.82954 | 0.8 | 0.88 | 0.87 | 0.71 | 0.84 | 0.9 | 0.85 | 0.82 | 0.82 | 0.84 | 0.82 | 0.86868 | 0.789 | 0.86 | 0.92 | 0.83 | 0.85 | 1.0107 | 0.9181 | 0.91485 |  |  |  |  | 16
| 1 | 4 | 0.9 | 0.8294 | 0.8 | 0.89 | 0.9 | 0.79 | 0.82 | 0.9 | 0.9 | 0.94 | 0.83 | 0.83 | 0.86 | 0.90753 | 0.807 | 0.84 | 0.89 | 0.78 | 0.81 | 0.99836 | 0.8742 | 0.91589 |  |  |  |  | 0.85987
| 1 | 5 | 0.92 | 0.84175 | 0.9 | 0.89 | 0.86 | 0.73 | 0.9 | 0.9 | 0.91 | 0.87 | 0.87 | 0.83 | 0.82 | 0.88928 | 0.793 | 0.8 | 0.88 | 0.79 | 0.82 | 1.0204 | 0.8754 | 0.90535 |  |  |  |  | 62
| 1 | 6 | 0.9 | 0.7842 | 0.8 | 0.88 | 0.89 | 0.78 | 0.84 | 0.9 | 0.87 | 0.84 | 0.88 | 0.85 | 0.88 | 0.857 | 0.846 | 0.78 | 0.87 | 0.87 | 0.82 | 1.0078 | 0.8986 | 0.91536
| Mean |  | 0.905 | 0.82127 | 0.81667 | 0.88333 | 0.87667 | 0.74333 | 0.85333 | 0.91667 | 0.89 | 0.85667 | 0.84667 | 0.83667 | 0.85333 | 0.88187 | 0.80983 | 0.825 | 0.89 | 0.82833 | 0.83167 | 1.0173 | 0.88732 | 0.90987
| Median |  | 0.905 | 0.82617 | 0.8 | 0.88 | 0.875 | 0.755 | 0.84 | 0.9 | 0.895 | 0.84 | 0.845 | 0.835 | 0.86 | 0.88436 | 0.8 | 0.83 | 0.89 | 0.815 | 0.825 | 1.0155 | 0.887 | 0.9118
| Std Dev. |  | 0.010488 | 0.019665 | 0.040825 | 0.005164 | 0.01633 | 0.048028 | 0.030111 | 0.040825 | 0.026077 | 0.044121 | 0.026583 | 0.008165 | 0.027325 | 0.017544 | 0.025663 | 0.030822 | 0.016733 | 0.047924 | 0.019408 | 0.01454 | 0.023724 | 0.0067878
| Rel.Std.Dev. |  | 1.159% | 2.394% | 4.999% | 0.5846% | 1.863% | 6.461% | 3.529% | 4.454% | 2.93% | 5.15% | 3.14% | 0.9759% | 3.202% | 1.989% | 3.169% | 3.736% | 1.88% | 5.786% | 2.334% | 1.429% | 2.674% | 0.746%
| PDM3 |  | 5.249% | -4.489% | -5.024% | 2.729% | 1.954% | -13.55% | -0.7597% | 6.606% | 3.505% | -0.372% | -1.535% | -2.698% | -0.7597% | 2.559% | -5.819% | -4.055% | 3.505% | -3.667% | -3.279% | 18.31% | 3.193% | 5.816%
| Z-Score (Absolute) |  | 1.16 | 0.78 | NA | 0.66 | 0.5 | 2.58 | 0.04 | NA | 0.81 | 0.04 | 0.19 | 0.42 | 0.04 | 0.62 | 1.04 | 0.69 | 0.81 | 0.62 | 0.54 | 3.75 | 0.75 | 1.27
Results from laboratories 1.05 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 235b (Certified Value 0.3 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.20; 1.08: 1.26; 1.09: 1.28; 1.10: 1.30; 1.11: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.22628; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.215; 1.08: 0.32; 1.09: 0.2; 1.10: 0.2394; 1.11: 0.42326 |  | 1
| 1 | 2 | 1.01: 0.21776; 1.03: 0.3; 1.05: 0.3; 1.06: 0.5; 1.07: 0.223; 1.08: 0.31; 1.09: 0.2; 1.10: 0.2258; 1.11: 0.43202 |  | 13
| 1 | 3 | 1.01: 0.22915; 1.03: 0.3; 1.05: 0.3; 1.06: 0.4; 1.07: 0.246; 1.08: 0.34; 1.09: 0.2; 1.10: 0.2196; 1.11: 0.4521 |  | 16
| 1 | 4 | 1.01: 0.21858; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.223; 1.08: 0.3; 1.09: 0.2; 1.10: 0.2236; 1.11: 0.42236 |  | 0.29508
| 1 | 5 | 1.01: 0.22541; 1.03: 0.35; 1.05: 0.3; 1.06: 0.3; 1.07: 0.222; 1.08: 0.31; 1.09: 0.2; 1.10: 0.2249; 1.11: 0.41785 |  | 63
| 1 | 6 | 1.01: 0.22844; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.22; 1.08: 0.31; 1.09: 0.2; 1.10: 0.2229; 1.11: 0.4304
| Mean |  | 1.01: 0.22427; 1.03: 0.30833; 1.05: 0.3; 1.06: 0.35; 1.07: 0.22483; 1.08: 0.315; 1.09: 0.2; 1.10: 0.22603; 1.11: 0.42967
| Median |  | 1.01: 0.22585; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.2225; 1.08: 0.31; 1.09: 0.2; 1.10: 0.22425; 1.11: 0.42683
| Std Dev. |  | 1.01: 0.0049241; 1.03: 0.020412; 1.05: 0; 1.06: 0.083666; 1.07: 0.010797; 1.08: 0.013784; 1.09: 3.0405e-17; 1.10: 0.0068867; 1.11: 0.012193
| Rel.Std.Dev. |  | 1.01: 2.196%; 1.03: 6.62%; 1.05: 0%; 1.06: 23.9%; 1.07: 4.802%; 1.08: 4.376%; 1.09: 1.52e-14%; 1.10: 3.047%; 1.11: 2.838%
| PDM3 |  | 1.01: -24%; 1.03: 4.491%; 1.05: 1.667%; 1.06: 18.61%; 1.07: -23.81%; 1.08: 6.75%; 1.09: -32.22%; 1.10: -23.4%; 1.11: 45.61%
| Z-Score (Absolute) |  | 1.01: 0.69; 1.03: 0; 1.05: NA; 1.06: 0.34; 1.07: 0.68; 1.08: 0.05; 1.09: NA; 1.10: 0.67; 1.11: 0.99
Results from laboratories 1.06 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 235b (Certified Value 2.81 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.87 | 2.7994 | 2.7 | 2.8 | 2.86 | 2.5 | 3.23 | 3 | 2.6 | 2.7 | 2.9 | 2.6 | 2.9 | 2.7641 | 2.8 | 2.488 | 2.69 | 2.92 | 3.1 | 2.7 | 5.4555 | 3.8202 | 3.1194 |  |  |  | 1
| 1 | 2 | 2.81 | 2.8045 | 2.8 | 2.9 | 2.91 | 2.5 | 3.25 | 3 | 2.7 | 2.7 | 2.8 | 2.6 | 3.1 | 2.7087 | 2.77 | 2.7 | 2.73 | 2.97 | 2.7 | 2.6 | 2.8195 | 3.9525 | 3.1462 |  |  |  | 35
| 1 | 3 | 2.76 | 2.8402 | 2.7 | 2.9 | 2.89 | 2.7 | 2.94 | 2.9 | 2.6 | 2.7 | 2.8 | 2.8 | 2.9 | 2.6798 | 2.71 | 2.56 | 2.83 | 2.97 | 2.9 | 2.7 | 2.8204 | 3.9321 | 3.1462 |  |  |  | 16
| 1 | 4 | 2.83 | 2.8035 | 2.6 | 2.9 | 2.9 | 2.8 | 3.07 | 2.9 | 2.6 | 2.9 | 2.9 | 2.7 | 3.1 | 2.8152 | 2.8 | 2.6 | 2.7 | 2.94 | 2.7 | 2.6 | 2.8096 | 3.9746 | 3.1211 |  |  |  | 2.8114
| 1 | 5 | 2.83 | 2.7923 | 2.8 | 3 | 2.93 | 2.8 | 2.87 | 2.9 | 2.6 | 2.9 | 2.8 | 2.7 | 2.7 | 2.7543 | 2.71 | 2.627 | 2.69 | 2.88 | 2.7 | 2.6 | 2.8533 | 3.9856 | 3.1183 |  |  |  | 64
| 1 | 6 | 2.86 | 2.8489 | 2.8 | 2.9 | 2.85 | 2.8 | 3.04 | 2.9 | 2.5 | 2.8 | 2.9 | 2.7 | 2.8 | 2.6481 | 3.5 | 2.797 | 2.62 | 2.84 | 3.1 | 2.6 | 2.7929 | 3.9037 | 3.1285
| Mean |  | 2.8267 | 2.8148 | 2.7333 | 2.9 | 2.89 | 2.6833 | 3.0667 | 2.9333 | 2.6 | 2.7833 | 2.85 | 2.6833 | 2.9167 | 2.7284 | 2.8817 | 2.6287 | 2.71 | 2.92 | 2.8667 | 2.6333 | 3.2585 | 3.9281 | 3.13
| Median |  | 2.83 | 2.804 | 2.75 | 2.9 | 2.895 | 2.75 | 3.055 | 2.9 | 2.6 | 2.75 | 2.85 | 2.7 | 2.9 | 2.7315 | 2.785 | 2.6135 | 2.695 | 2.93 | 2.8 | 2.6 | 2.8199 | 3.9423 | 3.1248
| Std Dev. |  | 0.039328 | 0.023583 | 0.08165 | 0.063246 | 0.030332 | 0.1472 | 0.15214 | 0.05164 | 0.063246 | 0.098319 | 0.054772 | 0.075277 | 0.16021 | 0.061134 | 0.30564 | 0.1084 | 0.068993 | 0.051769 | 0.19664 | 0.05164 | 1.0765 | 0.060527 | 0.013078
| Rel.Std.Dev. |  | 1.391% | 0.8378% | 2.987% | 2.181% | 1.05% | 5.486% | 4.961% | 1.76% | 2.433% | 3.532% | 1.922% | 2.805% | 5.493% | 2.241% | 10.61% | 4.124% | 2.546% | 1.773% | 6.859% | 1.961% | 33.04% | 1.541% | 0.4178%
| PDM3 |  | 0.5417% | 0.1195% | -2.778% | 3.15% | 2.794% | -4.557% | 9.078% | 4.336% | -7.521% | -0.9996% | 1.372% | -4.557% | 3.743% | -2.955% | 2.498% | -6.501% | -3.608% | 3.861% | 1.964% | -6.335% | 15.9% | 39.72% | 11.33%
| Z-Score (Absolute) |  | 0.13 | 0.2 | 0.67 | 0.29 | 0.23 | 0.96 | 1.25 | 0.48 | 1.44 | 0.39 | 0 | 0.96 | 0.39 | 0.7 | 0.18 | 1.28 | 0.81 | 0.4 | 0.1 | 1.25 | 2.36 | 6.23 | 1.62
| Analytical results for V in OREAS 235b (Certified Value 111 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 119 | 97.37 | 105 | 104 | 119.8715 | 114 | 111 | 106 | 110 | 110 | 109 | 105 | 107 | 111 | 112.0332 | 111 | 112 | 108.8 | 112 | 113 | 112 | 102.63 | 114 | 104.1841 | 112.674 |  | 1
| 1 | 2 | 116 | 96.29 | 110 | 96 | 120.0405 | 115 | 111 | 98 | 109 | 111 | 117 | 109 | 105 | 112 | 112.768 | 111 | 114 | 110.1 | 114 | 115 | 112 | 102.08 | 114 | 109.3319 | 114.5696 |  | 36
| 1 | 3 | 114 | 96.04 | 105 | 85 | 119.8855 | 115 | 111 | 104 | 109 | 107 | 112 | 107 | 107 | 111 | 115.2383 | 110 | 115 | 110.3 | 114 | 114 | 112 | 104.06 | 111 | 111.5037 | 111.6546 |  | 16
| 1 | 4 | 116 | 95.56 | 105 | 95 | 123.06 | 115 | 112 | 106 | 109 | 111 | 118 | 108 | 108 | 112 | 111.9542 | 112 | 113 | 109.6 | 114 | 115 | 112 | 101.97 | 111 | 108.0449 | 112.7012 |  | 110.9319
| 1 | 5 | 117 | 96.8 | 110 | 89 | 122.207 | 115 | 113 | 107 | 111 | 107 | 119 | 110 | 109 | 112 | 111.1394 | 109 | 114 | 109.1 | 113 | 114 | 109 | 102.96 | 108 | 110.6189 | 113.6778 |  | 65
| 1 | 6 | 119 | 96.49 | 110 | 85 | 122.769 | 115 | 114 | 105 | 108 | 108 | 115 | 111 | 110 | 114 | 112.9048 | 112 | 115 | 110.6 | 113 | 116 | 110 | 103.62 | 110 | 109.1308 | 114.4627
| Mean |  | 116.8333 | 96.425 | 107.5 | 92.3333 | 121.3056 | 114.8333 | 112 | 104.3333 | 109.3333 | 109 | 115 | 108.3333 | 107.6667 | 112 | 112.673 | 110.8333 | 113.8333 | 109.75 | 113.3333 | 114.5 | 111.1667 | 102.8867 | 111.3333 | 108.8024 | 113.29
| Median |  | 116.5 | 96.39 | 107.5 | 92 | 121.1238 | 115 | 111.5 | 105.5 | 109 | 109 | 116 | 108.5 | 107.5 | 112 | 112.4006 | 111 | 114 | 109.85 | 113.5 | 114.5 | 112 | 102.795 | 111 | 109.2314 | 113.1895
| Std Dev. |  | 1.9408 | 0.62491 | 2.7386 | 7.4207 | 1.5301 | 0.40825 | 1.2649 | 3.266 | 1.0328 | 1.8974 | 3.8471 | 2.1602 | 1.7512 | 1.0954 | 1.4086 | 1.169 | 1.169 | 0.7064 | 0.8165 | 1.0488 | 1.3292 | 0.83387 | 2.3381 | 2.5647 | 1.1457
| Rel.Std.Dev. |  | 1.661% | 0.6481% | 2.548% | 8.037% | 1.261% | 0.3555% | 1.129% | 3.13% | 0.9446% | 1.741% | 3.345% | 1.994% | 1.626% | 0.9781% | 1.25% | 1.055% | 1.027% | 0.6436% | 0.7204% | 0.916% | 1.196% | 0.8105% | 2.1% | 2.357% | 1.011%
| PDM3 |  | 5.32% | -13.08% | -3.094% | -16.77% | 9.351% | 3.517% | 0.9628% | -5.948% | -1.441% | -1.742% | 3.667% | -2.342% | -2.943% | 0.9628% | 1.569% | -0.08886% | 2.616% | -1.065% | 2.165% | 3.216% | 0.2116% | -7.252% | 0.3619% | -1.92% | 2.126%
| Z-Score (Absolute) |  | 1.43 | 3.73 | 0.93 | 4.76 | 2.56 | 0.93 | 0.21 | 1.73 | 0.46 | 0.55 | 0.97 | 0.72 | 0.89 | 0.21 | 0.38 | 0.08 | 0.67 | 0.36 | 0.55 | 0.84 | 0 | 2.09 | 0.04 | 0.6 | 0.54
| Analytical results for W in OREAS 235b (Certified Value 2.34 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 2.6 | 2.2842 | 1 | 0.2 | 20.9613 | 2.2 | 2 | 2.3 | 2.1 | 2.6 | 2.3 | 2.4 | 2.2 | 2.4 | 2.4208 | 3.4 | 2.1 | < 5 | 3.04 | 2.9 | 2 | 3.3556 | < 5 | 2.8851 |  |  | 1
| 1 | 2 | 2.4 | 2.2827 | 1 | < 0.1 | 18.8572 | 2.2 | 2 | 2.1 | 2.1 | 2.6 | 2.2 | 2.3 | 2.2 | 2.9 | 2.4016 | 3.3 | 2.2 | < 5 | 2.91 | 2.8 | 1.9 | 1.7022 | < 5 | 2.7744 |  |  | 37
| 1 | 3 | 2.4 | 2.225 | 1 | < 0.1 | 20.4923 | 2.2 | 2.1 | 1.7 | 2.2 | 2.5 | 2.2 | 2.3 | 2.3 | 2.7 | 2.4287 | 3.4 | 2.3 | < 5 | 2.9 | 2.7 | 1.9 | 1.6438 | < 5 | 2.8356 |  |  | 16
| 1 | 4 | 2.7 | 2.2507 | 1 | < 0.1 | 19.7735 | 2.3 | 2.1 | 1.8 | 2.5 | 2.8 | 2.6 | 2.3 | 2.2 | 2.9 | 2.3908 | 3.3 | 2.1 | < 5 | 2.88 | 2.8 | 1.8 | 1.6282 | < 5 | 2.915 |  |  | 2.3419
| 1 | 5 | 2.7 | 2.1813 | 1 | < 0.1 | 20.8125 | 2.3 | 2.2 | 2.1 | 2.6 | 2.4 | 2.4 | 2.5 | 2.2 | 2.5 | 2.361 | 3.4 | 2 | < 5 | 2.81 | 2.5 | 2 | 1.6121 | < 5 | 2.7527 |  |  | 66
| 1 | 6 | 2.7 | 2.2613 | 1 | < 0.1 | 20.3458 | 2.4 | 2.1 | 1.8 | 2.3 | 2.4 | 2.3 | 2.4 | 2.3 | 2.6 | 2.3702 | 3.4 | 2.3 | < 5 | 2.74 | 2.7 | 1.9 | 1.5964 | < 5 | 2.8002
| Mean |  | 2.5833 | 2.2475 | 1 | 0.2 | 20.2071 | 2.2667 | 2.0833 | 1.9667 | 2.3 | 2.55 | 2.3333 | 2.3667 | 2.2333 | 2.6667 | 2.3955 | 3.3667 | 2.1667 |  | 2.88 | 2.7333 | 1.9167 | 1.923 |  | 2.8272
| Median |  | 2.65 | 2.256 | 1 | 0.2 | 20.4191 | 2.25 | 2.1 | 1.95 | 2.25 | 2.55 | 2.3 | 2.35 | 2.2 | 2.65 | 2.3962 | 3.4 | 2.15 |  | 2.89 | 2.75 | 1.9 | 1.636 |  | 2.8179
| Std Dev. |  | 0.1472 | 0.039188 | 0 |  | 0.78054 | 0.08165 | 0.075277 | 0.23381 | 0.20976 | 0.15166 | 0.15055 | 0.08165 | 0.05164 | 0.20656 | 0.026955 | 0.05164 | 0.12111 |  | 0.10139 | 0.13663 | 0.075277 | 0.70275 |  | 0.063569
| Rel.Std.Dev. |  | 5.698% | 1.744% | 0% |  | 3.863% | 3.602% | 3.613% | 11.89% | 9.12% | 5.947% | 6.452% | 3.45% | 2.312% | 7.746% | 1.125% | 1.534% | 5.59% |  | 3.52% | 4.999% | 3.928% | 36.54% |  | 2.249%
| PDM3 |  | 10.31% | -4.029% | -57.3% | -91.46% | 762.9% | -3.211% | -11.04% | -16.02% | -1.787% | 8.888% | -0.3639% | 1.059% | -4.634% | 13.87% | 2.291% | 43.76% | -7.481% |  | 22.98% | 16.72% | -18.16% | -17.88% |  | 20.72%
| Z-Score (Absolute) |  | 0.67 | 0.37 | NA | 7.13 | 55.52 | 0.31 | 0.88 | 1.24 | 0.21 | 0.57 | 0.1 | 0 | 0.41 | 0.93 | 0.09 | 3.11 | 0.62 | 0.41 | 1.6 | 1.14 | 1.4 | 1.38 | 0.41 | 1.43
Results from laboratory 1.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 235b (Certified Value 16.4 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 16.92 | 13.6032 | 19.1 | 20.1 | 11.13 | 18.2 | 13 | 20.2 | 15.7 | 18.3 | 15.5 | 21.9 | 16.3 | 15 | 16.2237 | 21.8 | 12.31 | 14.5 | 19.6 | 17.5 | 15 | 12.68 | 14.909 | 25.7699 |  |  | 1
| 1 | 2 | 16.39 | 13.3425 | 19.5 | 20.4 | 10.96 | 17.8 | 10.6 | 19.7 | 14.9 | 18.7 | 16.5 | 16.8 | 16.6 | 15 | 15.9274 | 21.9 | 12.85 | 14.8 | 19.66 | 17 | 15.2 | 12.79 | 15.6018 | 26.2152 |  |  | 38
| 1 | 3 | 16.71 | 13.4872 | 18.7 | 20.5 | 10.535 | 17.8 | 12.4 | 18.7 | 14.3 | 17.9 | 16 | 16.8 | 17.3 | 15.4 | 15.98 | 21.2 | 13.91 | 15.1 | 19.21 | 17.1 | 15.2 | 13.13 | 15.7044 | 26.4815 |  |  | 16
| 1 | 4 | 17.35 | 13.0489 | 18.6 | 20 | 10.42 | 18.1 | 17.1 | 18.5 | 14.8 | 18.9 | 18.3 | 16.7 | 17.9 | 15.9 | 16.2436 | 21.3 | 12.94 | 14.6 | 19.54 | 17.8 | 14.8 | 12.71 | 15.3965 | 25.787 |  |  | 16.3526
| 1 | 5 | 18.14 | 13.2802 | 19.4 | 20.1 | 10.41 | 18.1 | 15.2 | 17.8 | 15.1 | 18.6 | 17.5 | 18 | 17.5 | 15.1 | 16.402 | 21.6 | 13.04 | 15.4 | 18.85 | 16.4 | 14.6 | 12.87 | 14.7808 | 25.4442 |  |  | 67
| 1 | 6 | 17.76 | 13.5139 | 19.4 | 19.6 | 10.165 | 17.9 | 14.7 | 18.1 | 14 | 18 | 17.2 | 18.6 | 20.8 | 15.6 | 16.187 | 20.9 | 12.62 | 14.5 | 19.08 | 17.7 | 15.2 | 12.89 | 15.8584 | 25.6048
| Mean |  | 17.2117 | 13.3793 | 19.1167 | 20.1167 | 10.6033 | 17.9833 | 13.8333 | 18.8333 | 14.8 | 18.4 | 16.8333 | 18.1333 | 17.7333 | 15.3333 | 16.1606 | 21.45 | 12.945 | 14.8167 | 19.3233 | 17.25 | 15 | 12.845 | 15.3751 | 25.8838
| Median |  | 17.135 | 13.4148 | 19.25 | 20.1 | 10.4775 | 18 | 13.85 | 18.6 | 14.85 | 18.45 | 16.85 | 17.4 | 17.4 | 15.25 | 16.2054 | 21.45 | 12.895 | 14.7 | 19.375 | 17.3 | 15.1 | 12.83 | 15.4992 | 25.7784
| Std Dev. |  | 0.66234 | 0.20012 | 0.38687 | 0.31885 | 0.36671 | 0.17224 | 2.3019 | 0.9331 | 0.6 | 0.4 | 1.0309 | 2.0017 | 1.6133 | 0.36697 | 0.1771 | 0.38341 | 0.54003 | 0.3656 | 0.32648 | 0.5244 | 0.25298 | 0.1627 | 0.43916 | 0.39
| Rel.Std.Dev. |  | 3.848% | 1.496% | 2.024% | 1.585% | 3.458% | 0.9578% | 16.64% | 4.954% | 4.054% | 2.174% | 6.124% | 11.04% | 9.097% | 2.393% | 1.096% | 1.787% | 4.172% | 2.468% | 1.69% | 3.04% | 1.687% | 1.267% | 2.856% | 1.507%
| PDM3 |  | 5.253% | -18.18% | 16.9% | 23.02% | -35.16% | 9.972% | -15.41% | 15.17% | -9.495% | 12.52% | 2.94% | 10.89% | 8.443% | -6.233% | -1.174% | 31.17% | -20.84% | -9.393% | 18.17% | 5.488% | -8.272% | -21.45% | -5.977% | 58.29%
| Z-Score (Absolute) |  | 0.06 | 1.19 | 0.69 | 1.01 | 2.1 | 0.31 | 1.04 | 0.59 | 0.73 | 0.45 | 0.06 | 0.36 | 0.23 | 0.55 | 0.28 | 1.45 | 1.34 | 0.72 | 0.75 | 0.07 | 0.66 | 1.37 | 0.54 | 2.9
| Analytical results for Yb in OREAS 235b (Certified Value 1.8 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.21; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.32; 1.14: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*OES; 1.13: 4A*OES; 1.14: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 1.5638; 1.03: 2.1; 1.05: 2.2; 1.06: 1.8; 1.07: 2.18; 1.08: 1.535; 1.09: 1.6; 1.10: 2.22; 1.11: 1.6; 1.12: 1.38; 1.13: 2.5344; 1.14: 2.8016 |  | 1
| 1 | 2 | 1.01: 1.5618; 1.03: 2.15; 1.05: 2.2; 1.06: 2.2; 1.07: 2.19; 1.08: 1.593; 1.09: 1.6; 1.10: 2.22; 1.11: 1.6; 1.12: 1.39; 1.13: 1.7698; 1.14: 2.9514 |  | 39
| 1 | 3 | 1.01: 1.6186; 1.03: 2.05; 1.05: 2.2; 1.06: 2.3; 1.07: 2.17; 1.08: 1.741; 1.09: 1.7; 1.10: 2.49; 1.11: 1.6; 1.12: 0.73; 1.13: 1.7698; 1.14: 2.9667 |  | 16
| 1 | 4 | 1.01: 1.5558; 1.03: 2; 1.05: 2.2; 1.06: 2.1; 1.07: 2.17; 1.08: 2.069; 1.09: 1.6; 1.10: 2.2; 1.11: 1.5; 1.12: 0.68; 1.13: 1.6848; 1.14: 2.7797 |  | 1.8043
| 1 | 5 | 1.01: 1.5866; 1.03: 2.1; 1.05: 2.2; 1.06: 2.1; 1.07: 2.14; 1.08: 1.601; 1.09: 1.6; 1.10: 2.17; 1.11: 1.6; 1.12: 0.69; 1.13: 1.6423; 1.14: 2.8027 |  | 68
| 1 | 6 | 1.01: 1.6465; 1.03: 2.15; 1.05: 2.2; 1.06: 1.8; 1.07: 2.13; 1.08: 1.597; 1.09: 1.5; 1.10: 2.12; 1.11: 1.6; 1.12: 1.39; 1.13: 1.7698; 1.14: 2.8339
| Mean |  | 1.01: 1.5888; 1.03: 2.0917; 1.05: 2.2; 1.06: 2.05; 1.07: 2.1633; 1.08: 1.6893; 1.09: 1.6; 1.10: 2.2367; 1.11: 1.5833; 1.12: 1.0433; 1.13: 1.8618; 1.14: 2.856
| Median |  | 1.01: 1.5752; 1.03: 2.1; 1.05: 2.2; 1.06: 2.1; 1.07: 2.17; 1.08: 1.599; 1.09: 1.6; 1.10: 2.21; 1.11: 1.6; 1.12: 1.055; 1.13: 1.7698; 1.14: 2.8183
| Std Dev. |  | 1.01: 0.036496; 1.03: 0.058452; 1.05: 4.8648e-16; 1.06: 0.20736; 1.07: 0.023381; 1.08: 0.19811; 1.09: 0.063246; 1.10: 0.12972; 1.11: 0.040825; 1.12: 0.37649; 1.13: 0.33385; 1.14: 0.081807
| Rel.Std.Dev. |  | 1.01: 2.297%; 1.03: 2.795%; 1.05: 2.211e-14%; 1.06: 10.12%; 1.07: 1.081%; 1.08: 11.73%; 1.09: 3.953%; 1.10: 5.8%; 1.11: 2.578%; 1.12: 36.09%; 1.13: 17.93%; 1.14: 2.864%
| PDM3 |  | 1.01: -11.94%; 1.03: 15.93%; 1.05: 21.93%; 1.06: 13.62%; 1.07: 19.9%; 1.08: -6.371%; 1.09: -11.32%; 1.10: 23.96%; 1.11: -12.25%; 1.12: -42.17%; 1.13: 3.188%; 1.14: 58.29%
| Z-Score (Absolute) |  | 1.01: 0.9; 1.03: 0.33; 1.05: 0.6; 1.06: 0.23; 1.07: 0.51; 1.08: 0.66; 1.09: 0.88; 1.10: 0.69; 1.11: 0.92; 1.12: 2.25; 1.13: 0.23; 1.14: 2.22
| Analytical results for Zn in OREAS 235b (Certified Value 97 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 101 | 91.0096 | 94 | 101 | 96.48 | 99 | 101 | 98 | 98 | 97 | 93 | 92 | 96 | 97.7 | 96.1619 | 103 | 95.16 | 96 | 110.3 | 94 | 94 | 95.5 | 99.165 | 104 | 86.725 | 95.6077 | 1
| 1 | 2 | 99 | 92.3408 | 94 | 102 | 99.3788 | 97 | 100 | 94 | 99 | 97 | 100 | 92 | 95 | 97.5 | 96.3379 | 103 | 102.3 | 100 | 105.9 | 94 | 96 | 99.4 | 98.164 | 102 | 84.8854 | 95.008 | 40
| 1 | 3 | 100 | 93.4328 | 92 | 104 | 98.1627 | 97 | 99 | 124 | 98 | 95 | 96 | 92 | 98 | 97.4 | 97.2929 | 104 | 94.81 | 97 | 109.5 | 94 | 95 | 92.9 | 99.583 | 99 | 85.9863 | 93.6475 | 16
| 1 | 4 | 101 | 89.4392 | 90 | 103 | 98.2886 | 100 | 104 | 95 | 98 | 97 | 107 | 93 | 96 | 98.2 | 95.1239 | 102 | 100.7 | 97 | 104.6 | 95 | 94 | 92.3 | 97.559 | 100 | 87.1741 | 94.4158 | 97.2547
| 1 | 5 | 100 | 89.9592 | 96 | 104 | 96.8836 | 99 | 103 | 96 | 101 | 93 | 103 | 95 | 97 | 98.9 | 96.5628 | 102 | 99.13 | 99 | 111.1 | 94 | 96 | 91.2 | 98.769 | 98 | 83.799 | 94.6142 | 69
| 1 | 6 | 104 | 93.0792 | 94 | 105 | 94.9254 | 100 | 98 | 94 | 100 | 94 | 100 | 95 | 96 | 100 | 96.2631 | 101 | 101.8 | 96 | 106.9 | 94 | 96 | 91.2 | 98.593 | 99 | 83.0024 | 94.7039
| Mean |  | 100.8333 | 91.5435 | 93.3333 | 103.1667 | 97.3532 | 98.6667 | 100.8333 | 100.1667 | 99 | 95.5 | 99.8333 | 93.1667 | 96.3333 | 98.2833 | 96.2904 | 102.5 | 98.9833 | 97.5 | 108.05 | 94.1667 | 95.1667 | 93.75 | 98.6388 | 100.3333 | 85.262 | 94.6662
| Median |  | 100.5 | 91.6752 | 94 | 103.5 | 97.5231 | 99 | 100.5 | 95.5 | 98.5 | 96 | 100 | 92.5 | 96 | 97.95 | 96.3005 | 102.5 | 99.915 | 97 | 108.2 | 94 | 95.5 | 92.6 | 98.681 | 99.5 | 85.4359 | 94.6591
| Std Dev. |  | 1.7224 | 1.6605 | 2.0656 | 1.472 | 1.5818 | 1.3663 | 2.3166 | 11.7714 | 1.2649 | 1.7607 | 4.9565 | 1.472 | 1.0328 | 1.0068 | 0.70135 | 1.0488 | 3.2844 | 1.6432 | 2.6197 | 0.40825 | 0.98319 | 3.1879 | 0.71817 | 2.2509 | 1.6558 | 0.64928
| Rel.Std.Dev. |  | 1.708% | 1.814% | 2.213% | 1.427% | 1.625% | 1.385% | 2.297% | 11.75% | 1.278% | 1.844% | 4.965% | 1.58% | 1.072% | 1.024% | 0.7284% | 1.023% | 3.318% | 1.685% | 2.425% | 0.4335% | 1.033% | 3.4% | 0.7281% | 2.243% | 1.942% | 0.6859%
| PDM3 |  | 3.68% | -5.872% | -4.032% | 6.079% | 0.1013% | 1.452% | 3.68% | 2.994% | 1.795% | -1.804% | 2.651% | -4.203% | -0.9473% | 1.058% | -0.9915% | 5.393% | 1.777% | 0.2523% | 11.1% | -3.175% | -2.147% | -3.604% | 1.423% | 3.166% | -12.33% | -2.662%
| Z-Score (Absolute) |  | 0.77 | 1.66 | 1.19 | 1.38 | 0.14 | 0.2 | 0.77 | 0.59 | 0.29 | 0.62 | 0.51 | 1.23 | 0.41 | 0.1 | 0.42 | 1.2 | 0.28 | 0.1 | 2.65 | 0.97 | 0.71 | 1.08 | 0.2 | 0.64 | 3.3 | 0.84
| Analytical results for Zr in OREAS 235b (Certified Value 140 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 147.1 | 135.3702 | 142 | 151 | 110 | 132.2 | 136.5 | 161 | 142.9 | 141.5 | 135 | 141.5 | 141 | 132.4 | 137.7996 | 147 | 136 | 149.4 | 144 | 134.2 | 132.9 | 123.431 | 152 | 143.3186 | 225.0713 |  | 1
| 1 | 2 | 146.3 | 135.0528 | 143 | 124 | 112.1158 | 135.3 | 133.4 | 153 | 148 | 141 | 139.5 | 136.5 | 139 | 131.5 | 138.0056 | 149 | 136 | 150.7 | 144.3 | 129.3 | 129.2 | 121.858 | 149 | 138.3335 | 224.9821 |  | 41
| 1 | 3 | 147.6 | 140.2723 | 146 | 104 | 111.39 | 133.3 | 143.6 | 141 | 142.4 | 133.5 | 135.5 | 140 | 143.5 | 132.5 | 135.4771 | 150 | 135 | 151.6 | 140.3 | 132 | 134.2 | 124.982 | 144 | 138.0817 | 230.4506 |  | 16
| 1 | 4 | 151.4 | 141.4365 | 142 | 98 | 114.52 | 138.6 | 138.7 | 145 | 141.5 | 145.5 | 157.5 | 137 | 145.5 | 133.8 | 136.949 | 146 | 134 | 147.4 | 141.7 | 133.9 | 131.3 | 123.123 | 145 | 137.3264 | 227.4446 |  | 139.6935
| 1 | 5 | 152.9 | 139.7286 | 140 | 102 | 112.54 | 140 | 131.1 | 141 | 145.1 | 137.5 | 150 | 143.5 | 141.5 | 134.9 | 136.8449 | 147 | 135 | 149 | 140 | 127.3 | 126.1 | 122.199 | 145 | 128.8164 | 225.9874 |  | 70
| 1 | 6 | 152.7 | 136.6895 | 147 | 101 | 114.54 | 138.3 | 128.8 | 143 | 143.9 | 139 | 143.5 | 144 | 141 | 135.1 | 136.3891 | 146 | 134 | 148.8 | 136.6 | 133.5 | 128.1 | 122.65 | 141 | 139.3406 | 226.2743
| Mean |  | 149.6667 | 138.0916 | 143.3333 | 113.3333 | 112.5176 | 136.2833 | 135.35 | 147.3333 | 143.9667 | 139.6667 | 143.5 | 140.4167 | 141.9167 | 133.3667 | 136.9109 | 147.5 | 135 | 149.4833 | 141.15 | 131.7 | 130.3 | 123.0405 | 146 | 137.5362 | 226.7017
| Median |  | 149.5 | 138.209 | 142.5 | 103 | 112.3279 | 136.8 | 134.95 | 144 | 143.4 | 140 | 141.5 | 140.75 | 141.25 | 133.15 | 136.8969 | 147 | 135 | 149.2 | 141 | 132.75 | 130.25 | 122.8865 | 145 | 138.2076 | 226.1309
| Std Dev. |  | 2.9951 | 2.7287 | 2.6583 | 20.6656 | 1.7817 | 3.1556 | 5.3921 | 8.0416 | 2.3338 | 4.0579 | 8.8374 | 3.1846 | 2.2675 | 1.4638 | 0.92974 | 1.6432 | 0.89443 | 1.4838 | 2.8669 | 2.8121 | 3.0548 | 1.1124 | 3.8987 | 4.7687 | 2.0452
| Rel.Std.Dev. |  | 2.001% | 1.976% | 1.855% | 18.23% | 1.583% | 2.315% | 3.984% | 5.458% | 1.621% | 2.905% | 6.158% | 2.268% | 1.598% | 1.098% | 0.6791% | 1.114% | 0.6625% | 0.9926% | 2.031% | 2.135% | 2.344% | 0.9041% | 2.67% | 3.467% | 0.9022%
| PDM3 |  | 7.139% | -1.147% | 2.606% | -18.87% | -19.45% | -2.441% | -3.109% | 5.469% | 3.059% | -0.01919% | 2.725% | 0.5177% | 1.591% | -4.529% | -1.992% | 5.588% | -3.36% | 7.008% | 1.043% | -5.722% | -6.724% | -11.92% | 4.515% | -1.544% | 62.29%
| Z-Score (Absolute) |  | 1.44 | 0.23 | 0.53 | 3.81 | 3.92 | 0.49 | 0.62 | 1.11 | 0.62 | 0 | 0.55 | 0.11 | 0.33 | 0.91 | 0.4 | 1.13 | 0.67 | 1.42 | 0.21 | 1.15 | 1.35 | 2.4 | 0.92 | 0.31 | 12.58
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C638: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C801: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C819: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C837: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C855: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1000: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1055: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1092: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1166: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1203: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 235b (Certified Value 0.145 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.14 | 0.163 | < 0.5 | 0.14 | 0.12 | 0.16 | 0.1 | 0.14 | 0.15 | 0.15 | 0.14 | 0.116 | < 1 | 0.3 | 0.16 | 0.23 | < 0.5 | 0.14 | 0.17 | < 1 | 1.5066 |  |  |  | 1
| 1 | 2 | 0.13 | 0.157 | < 0.5 | 0.15 | 0.1 | 0.17 | 0.1 | 0.14 | 0.15 | 0.14 | 0.15 | 0.128 | < 1 | 0.3 | 0.15 | 0.13 | < 0.5 | 0.16 | 0.15 | < 1 | 1.8019 |  |  |  | 25
| 1 | 3 | 0.14 | 0.147 | < 0.5 | 0.16 | 0.11 | 0.18 | 0.1 | 0.14 | 0.15 | 0.14 | 0.15 | 0.119 | < 1 | 0.3 | 0.14 | 0.11 | < 0.5 | 0.15 | 0.14 | < 1 | 1.7009 |  |  |  | 16
| 1 | 4 | 0.14 | 0.172 | < 0.5 | 0.14 | 0.11 | 0.17 | 0.1 | 0.15 | 0.15 | 0.15 | 0.15 | 0.14 | < 1 | 0.2 | 0.16 | 0.15 | < 0.5 | 0.13 | 0.13 | < 1 | 1.2119 |  |  |  | 0.14502
| 1 | 5 | 0.13 | 0.154 | < 0.5 | 0.14 | 0.1 | 0.16 | 0.1 | 0.14 | 0.15 | 0.15 | 0.15 | 0.126 | < 1 | 0.2 | 0.14 | 0.12 | < 0.5 | 0.15 | 0.16 | < 1 | 1.4524 |  |  |  | 72
| 1 | 6 | 0.14 | 0.149 | < 0.5 | 0.14 | 0.11 | 0.16 | 0.1 | 0.14 | 0.15 | 0.14 | 0.15 | 0.13 | < 1 | < 0.2 | 0.15 | 0.1 | < 0.5 | 0.14 | 0.15 | < 1 | 1.1277
| Mean |  | 0.13667 | 0.157 |  | 0.145 | 0.10833 | 0.16667 | 0.1 | 0.14167 | 0.15 | 0.145 | 0.14833 | 0.1265 |  | 0.26 | 0.15 | 0.14 |  | 0.145 | 0.15 |  | 1.4669
| Median |  | 0.14 | 0.1555 |  | 0.14 | 0.11 | 0.165 | 0.1 | 0.14 | 0.15 | 0.145 | 0.15 | 0.127 |  | 0.3 | 0.15 | 0.125 |  | 0.145 | 0.15 |  | 1.4795
| Std Dev. |  | 0.005164 | 0.0093167 |  | 0.0083666 | 0.0075277 | 0.008165 | 1.5202e-17 | 0.0040825 | 0 | 0.0054772 | 0.0040825 | 0.0085264 |  | 0.054772 | 0.0089443 | 0.047329 |  | 0.010488 | 0.014142 |  | 0.26414
| Rel.Std.Dev. |  | 3.779% | 5.934% |  | 5.77% | 6.949% | 4.899% | 1.52e-14% | 2.882% | 0% | 3.777% | 2.752% | 6.74% |  | 21.07% | 5.963% | 33.81% |  | 7.233% | 9.428% |  | 18.01%
| PDM3 |  | -5.757% | 8.265% |  | -0.01045% | -25.3% | 14.93% | -31.04% | -2.309% | 3.437% | -0.01045% | 2.288% | -12.77% |  | 79.29% | 3.437% | -3.458% |  | -0.01045% | 3.437% |  | 911.5%
| Z-Score (Absolute) |  | 0.9 | 0.47 | 6.74 | 0.34 | 2.81 | 1.12 | NA | 0.56 | 0 | 0.34 | 0.11 | 1.58 | 23.6 | NA | 0 | 0.67 | 6.74 | 0.34 | 0 | 23.6 | 88.8
Results from laboratories 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 235b (Certified Value 3.02 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 2.785 | 3.5164 | 3.05 | 3.2567 | 3.272 | 3.1 | 3.53 | 2.91 | 2.83 | 3.17 | 3 | 3.04 | 2.99 | 3.0638 | 2.805 | 2.99 | 3.03 | 2.98 | 2.7495 | 4.9866 |  |  |  |  | 1
| 1 | 2 | 2.774 | 3.5013 | 3.12 | 3.2767 | 3.255 | 2.98 | 3.59 | 2.93 | 2.92 | 3.13 | 3.02 | 2.98 | 3.13 | 2.9875 | 2.752 | 3.2 | 3.04 | 2.99 | 2.716 | 4.9784 |  |  |  |  | 0
| 1 | 3 | 2.815 | 3.5025 | 2.91 | 3.31 | 3.177 | 3 | 3.51 | 2.91 | 2.91 | 3.1 | 3.06 | 3.03 | 3.1 | 3.0786 | 2.826 | 3.02 | 3.09 | 2.86 | 2.6939 | 4.9762 |  |  |  |  | 16
| 1 | 4 | 2.786 | 3.4907 | 3.13 | 3.25 | 3.172 | 2.98 | 3.72 | 2.88 | 2.96 | 3.07 | 3.06 | 3.04 | 2.99 | 3.0155 | 2.789 | 3.03 | 3.01 | 2.83 | 2.6844 | 4.9775 |  |  |  |  | 3.0237
| 1 | 5 | 2.802 | 3.4931 | 2.97 | 3.2667 | 3.188 | 3.03 | 3.84 | 2.89 | 2.91 | 3.12 | 3.03 | 3.08 | 2.93 | 3.1906 | 2.879 | 2.84 | 2.98 | 2.83 | 2.7362 | 4.9801 |  |  |  |  | 73
| 1 | 6 | 2.771 | 3.5195 | 3.05 | 3.31 | 3.158 | 3.12 | 3.65 | 2.94 | 2.87 | 3.12 | 3.06 | 3.06 | 3.14 | 3.0057 | 2.821 | 2.91 | 3.05 | 2.91 | 2.7783 | 4.9768
| Mean |  | 2.7888 | 3.5039 | 3.0383 | 3.2783 | 3.2037 | 3.035 | 3.64 | 2.91 | 2.9 | 3.1183 | 3.0383 | 3.0383 | 3.0467 | 3.057 | 2.812 | 2.9983 | 3.0333 | 2.9 | 2.7264 | 4.9793
| Median |  | 2.7855 | 3.5019 | 3.05 | 3.2717 | 3.1825 | 3.015 | 3.62 | 2.91 | 2.91 | 3.12 | 3.045 | 3.04 | 3.045 | 3.0397 | 2.813 | 3.005 | 3.035 | 2.885 | 2.7261 | 4.9779
| Std Dev. |  | 0.016845 | 0.011825 | 0.085421 | 0.026141 | 0.047643 | 0.061237 | 0.1249 | 0.022804 | 0.044721 | 0.033116 | 0.025626 | 0.033714 | 0.087788 | 0.074226 | 0.042294 | 0.12254 | 0.037238 | 0.072111 | 0.03536 | 0.0038427
| Rel.Std.Dev. |  | 0.604% | 0.3375% | 2.811% | 0.7974% | 1.487% | 2.018% | 3.431% | 0.7836% | 1.542% | 1.062% | 0.8434% | 1.11% | 2.881% | 2.428% | 1.504% | 4.087% | 1.228% | 2.487% | 1.297% | 0.07717%
| PDM3 |  | -7.768% | 15.88% | 0.4836% | 8.421% | 5.951% | 0.3733% | 20.38% | -3.761% | -4.091% | 3.129% | 0.4836% | 0.4836% | 0.7592% | 1.1% | -7.002% | -0.8393% | 0.3182% | -4.091% | -9.833% | 64.67%
| Z-Score (Absolute) |  | 1.26 | 2.35 | 0 | 1.21 | 0.84 | 0.02 | 3.04 | 0.65 | 0.7 | 0.4 | 0 | 0 | 0.04 | 0.09 | 1.14 | 0.2 | 0.03 | 0.7 | 1.58 | 9.82
| Analytical results for As in OREAS 235b (Certified Value 352 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 362.78 | 371.13 | 331 | 352.665 | 358 | 362.5 | 346 | 357.1 | 350 | 379 | 367 | 361 | 359.8 | 365.1766 | 337 | 332 | 392.87 | 362.9 | 352.8 | 328.5 | 341 | 680.5848 | 355.1232 |  | 1
| 1 | 2 | 362.79 | 369.17 | 330 | 352.1604 | 354 | 367 | 342 | 349 | 357 | 372 | 362 | 359 | 369.7 | 352.1568 | 339 | 319 | 349.25 | 362.9 | 340.7 | 320.4 | 346 | 689.7807 | 348.5805 |  | 26
| 1 | 3 | 365.49 | 375.86 | 322 | 347.3496 | 336 | 363.8 | 335 | 346.5 | 355 | 377 | 365 | 370 | 352.1 | 354.7318 | 337 | 334 | 362.51 | 360.6 | 339 | 319.1 | 350 | 541.3449 | 347.5589 |  | 16
| 1 | 4 | 366.69 | 373.6 | 333 | 346.88 | 344 | 361.7 | 337 | 345.1 | 364 | 375 | 375 | 362 | 398.6 | 354.9553 | 335 | 330 | 394.58 | 363.8 | 326.1 | 308.6 | 349 | 508.3628 | 351.4853 |  | 352.156
| 1 | 5 | 360.66 | 367.64 | 319 | 352.335 | 340 | 361.6 | 349 | 347.2 | 356 | 369 | 364 | 376 | 354.4 | 347.3607 | 340 | 310 | 373.29 | 367.9 | 328.9 | 328.6 | 359 | 638.9125 | 346.5138 |  | 74
| 1 | 6 | 359.63 | 367.93 | 320 | 350.5269 | 347 | 360.7 | 345 | 355.6 | 355 | 376 | 364 | 368 | 366.5 | 352.9968 | 334 | 320 | 376.85 | 364.8 | 330.8 | 327.9 | 368 | 673.5487 | 344.952
| Mean |  | 363.0067 | 370.8883 | 325.8333 | 350.3195 | 346.5 | 362.8833 | 342.3333 | 350.0833 | 356.1667 | 374.6667 | 366.1667 | 366 | 366.85 | 354.563 | 337 | 324.1667 | 374.8917 | 363.8167 | 336.3833 | 322.1833 | 352.1667 | 622.0891 | 349.0356
| Median |  | 362.785 | 370.15 | 326 | 351.3436 | 345.5 | 362.1 | 343.5 | 348.1 | 355.5 | 375.5 | 364.5 | 365 | 363.15 | 353.8643 | 337 | 325 | 375.07 | 363.35 | 334.9 | 324.15 | 349.5 | 656.2306 | 348.0697
| Std Dev. |  | 2.7108 | 3.298 | 6.1779 | 2.5941 | 8.3367 | 2.2693 | 5.4283 | 5.0364 | 4.535 | 3.6148 | 4.6224 | 6.4807 | 16.9622 | 5.8813 | 2.2804 | 9.347 | 17.4817 | 2.4359 | 9.8783 | 7.886 | 9.7451 | 77.9594 | 3.7009
| Rel.Std.Dev. |  | 0.7468% | 0.8892% | 1.896% | 0.7405% | 2.406% | 0.6253% | 1.586% | 1.439% | 1.273% | 0.9648% | 1.262% | 1.771% | 4.624% | 1.659% | 0.6767% | 2.883% | 4.663% | 0.6695% | 2.937% | 2.448% | 2.767% | 12.53% | 1.06%
| PDM3 |  | 3.081% | 5.319% | -7.475% | -0.5215% | -1.606% | 3.046% | -2.789% | -0.5886% | 1.139% | 6.392% | 3.979% | 3.931% | 4.173% | 0.6835% | -4.304% | -7.948% | 6.456% | 3.311% | -4.479% | -8.511% | 0.003028% | 76.65% | -0.8861%
| Z-Score (Absolute) |  | 0.49 | 0.95 | 1.67 | 0.25 | 0.47 | 0.48 | 0.71 | 0.26 | 0.09 | 1.17 | 0.67 | 0.66 | 0.71 | 0 | 1.02 | 1.77 | 1.18 | 0.54 | 1.06 | 1.88 | 0.14 | 15.55 | 0.32
| Analytical results for B in OREAS 235b (Indicative Value 15.9 ppm) | [25 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.06; 1.07: 1.07; 1.08: 1.08; 1.09: 1.12; 1.10: 1.13; 1.11: 1.14; 1.12: 1.15; 1.13: 1.16; 1.14: 1.27; 1.15: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES; 1.08: AR*OES/MS; 1.09: AR*MS; 1.10: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*OES/MS; 1.15: AR*OES |  |  |  | 1
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 01g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g; 1.14: 0.5g; 1.15: 0.5g |  |  |  | 1
| 1 | 1 | 1.01: < 5; 1.06: 4; 1.07: 26; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21 |  |  |  | 1
| 1 | 2 | 1.01: < 5; 1.06: 5; 1.07: 30; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21 |  |  |  | 4
| 1 | 3 | 1.01: < 5; 1.06: 5; 1.07: 24.588; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21 |  |  |  | 16
| 1 | 4 | 1.01: < 5; 1.06: 5; 1.07: 30; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21 |  |  |  | 15.9372
| 1 | 5 | 1.01: < 5; 1.06: 5; 1.07: 27.205; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 22 |  |  |  | 14
| 1 | 6 | 1.01: < 5; 1.06: 5; 1.07: 27.7; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 22
| Mean |  | 1.06: 4.8333; 1.07: 27.5822; 1.10: 10; 1.15: 21.3333
| Median |  | 1.06: 5; 1.07: 27.4525; 1.10: 10; 1.15: 21
| Std Dev. |  | 1.06: 0.40825; 1.07: 2.1595; 1.10: 0; 1.15: 0.5164
| Rel.Std.Dev. |  | 1.06: 8.447%; 1.07: 7.829%; 1.10: 0%; 1.15: 2.421%
| PDM3 |  | 1.06: -69.67%; 1.07: 73.07%; 1.10: -37.25%; 1.15: 33.86%
| Z-Score (Absolute) |  | 1.01: 10.11; 1.06: 0.67; 1.07: 91.36; 1.08: 0; 1.09: 0; 1.10: 20.23; 1.11: 0; 1.12: 0; 1.13: 20.23; 1.14: 0; 1.15: 66.08
| Analytical results for Ba in OREAS 235b (Certified Value 96 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 87.88 | 110.9746 | 111 | 101.3733 | 101 | 104 | 119 | 93 | 90 | 90 | 100 | 90 | 94.1 | 94.8112 | 90 | 93.64 | 79 | 100.7 | 105.4 | 101 | 94.86 | 604.642 |  |  | 1
| 1 | 2 | 86.47 | 108.8042 | 111 | 101.34 | 101 | 98 | 120 | 95 | 90 | 90 | 100 | 90 | 99 | 93.7519 | 89 | 91.26 | 102 | 100.97 | 101.2 | 101 | 93.18 | 608.4901 |  |  | 27
| 1 | 3 | 88.72 | 105.4684 | 109 | 102.2833 | 98 | 101 | 113 | 91 | 90 | 90 | 100 | 90 | 97.3 | 99.7606 | 89 | 101.6 | 79 | 98.54 | 105.6 | 94 | 92.83 | 601.2306 |  |  | 16
| 1 | 4 | 88.47 | 103.7369 | 114 | 100.54 | 98 | 102 | 123 | 92 | 90 | 90 | 100 | 90 | 93.8 | 92.977 | 90 | 96.34 | 80 | 100.57 | 103.5 | 94 | 91.64 | 625.03 |  |  | 96.0837
| 1 | 5 | 87.42 | 107.4817 | 111 | 103.42 | 97 | 107 | 119 | 93 | 90 | 90 | 100 | 100 | 92.6 | 93.6139 | 91 | 94.39 | 74 | 100.25 | 104.9 | 96 | 93.76 | 597.4158 |  |  | 75
| 1 | 6 | 88.29 | 105.1409 | 108 | 103.0067 | 102 | 106 | 123 | 93 | 90 | 90 | 100 | 90 | 99 | 95.8226 | 88 | 94.38 | 77 | 104.07 | 106.5 | 99 | 94.83 | 600.7585
| Mean |  | 87.875 | 106.9345 | 110.6667 | 101.9939 | 99.5 | 103 | 119.5 | 92.8333 | 90 | 90 | 100 | 91.6667 | 95.9667 | 95.1229 | 89.5 | 95.2683 | 81.8333 | 100.85 | 104.5167 | 97.5 | 93.5167 | 606.2612
| Median |  | 88.085 | 106.4751 | 111 | 101.8283 | 99.5 | 103 | 119.5 | 93 | 90 | 90 | 100 | 90 | 95.7 | 94.2816 | 89.5 | 94.385 | 79 | 100.635 | 105.15 | 97.5 | 93.47 | 602.9363
| Std Dev. |  | 0.82721 | 2.6724 | 2.0656 | 1.1018 | 2.0736 | 3.3466 | 3.6742 | 1.3292 | 0 | 0 | 0 | 4.0825 | 2.8176 | 2.4842 | 1.0488 | 3.5076 | 10.1078 | 1.7989 | 1.901 | 3.2711 | 1.2406 | 9.9367
| Rel.Std.Dev. |  | 0.9413% | 2.499% | 1.866% | 1.08% | 2.084% | 3.249% | 3.075% | 1.432% | 0% | 0% | 0% | 4.454% | 2.936% | 2.612% | 1.172% | 3.682% | 12.35% | 1.784% | 1.819% | 3.355% | 1.327% | 1.639%
| PDM3 |  | -8.543% | 11.29% | 15.18% | 6.151% | 3.556% | 7.198% | 24.37% | -3.383% | -6.332% | -6.332% | 4.076% | -4.597% | -0.1218% | -1% | -6.852% | -0.8486% | -14.83% | 4.961% | 8.777% | 1.474% | -2.672% | 531%
| Z-Score (Absolute) |  | 1.04 | 1.19 | 1.63 | 0.62 | 0.32 | 0.73 | 2.66 | 0.46 | 0.79 | 0.79 | 0.38 | 0.59 | 0.09 | 0.19 | 0.85 | 0.17 | 1.74 | 0.48 | 0.91 | 0.09 | 0.38 | 59.61
| Analytical results for Be in OREAS 235b (Certified Value 1.28 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 1.26 | 1.5025 | 1.3 | < 5 | 1.3 | 1.34 | 0.95 | 1 | 1.34 | 1.32 | 1.43 | 1.35 | 1.3794 | 1.4 | 1.241 | 1.5 | 1.28 | < 2 | 1.23 | 1.033 | 1.1237 | 1.3272 |  |  | 1
| 1 | 2 | 1.26 | 1.604 | 1.4 | < 5 | 1.3 | 1.27 | 0.91 | 1 | 1.38 | 1.34 | 1.37 | 1.35 | 1.4366 | 1.3 | 1.191 | 1.4 | 1.25 | < 2 | 1.18 | 1.017 | 1.1455 | 1.2981 |  |  | 28
| 1 | 3 | 1.29 | 1.7679 | 1.3 | < 5 | 1.3 | 1.24 | 1 | 1 | 1.42 | 1.3 | 1.38 | 1.34 | 1.3566 | 1.4 | 1.404 | 1.5 | 1.28 | < 2 | 1.16 | 1.014 | 1.1239 | 1.3362 |  |  | 16
| 1 | 4 | 1.24 | 1.8256 | 1.3 | < 5 | 1.3 | 1.28 | 0.94 | 1 | 1.5 | 1.28 | 1.37 | 1.38 | 1.3494 | 1.4 | 1.287 | 1.4 | 1.27 | < 2 | 1.17 | 1.011 | 1.059 | 1.3414 |  |  | 1.2774
| 1 | 5 | 1.22 | 1.5097 | 1.2 | < 5 | 1.3 | 1.29 | 0.91 | 1 | 1.42 | 1.34 | 1.45 | 1.4 | 1.4875 | 1.4 | 1.264 | 1.5 | 1.31 | < 2 | 1.11 | 1.03 | 1.1675 | 1.2298 |  |  | 76
| 1 | 6 | 1.3 | 1.582 | 1.3 | < 5 | 1.3 | 1.31 | 0.96 | 2 | 1.36 | 1.32 | 1.37 | 1.38 | 1.4586 | 1.3 | 1.266 | 1.6 | 1.3 | < 2 | 1.16 | 1.053 | 1.1459 | 1.3259
| Mean |  | 1.2617 | 1.632 | 1.3 |  | 1.3 | 1.2883 | 0.945 | 1.1667 | 1.4033 | 1.3167 | 1.395 | 1.3667 | 1.4113 | 1.3667 | 1.2755 | 1.4833 | 1.2817 |  | 1.1683 | 1.0263 | 1.1276 | 1.3098
| Median |  | 1.26 | 1.593 | 1.3 |  | 1.3 | 1.285 | 0.945 | 1 | 1.4 | 1.32 | 1.375 | 1.365 | 1.408 | 1.4 | 1.265 | 1.5 | 1.28 |  | 1.165 | 1.0235 | 1.1347 | 1.3265
| Std Dev. |  | 0.029944 | 0.13489 | 0.063246 |  | 0 | 0.034303 | 0.033912 | 0.40825 | 0.057155 | 0.023381 | 0.035637 | 0.023381 | 0.057465 | 0.05164 | 0.070989 | 0.075277 | 0.02137 |  | 0.038687 | 0.015769 | 0.037369 | 0.041946
| Rel.Std.Dev. |  | 2.373% | 8.266% | 4.865% |  | 0% | 2.663% | 3.589% | 34.99% | 4.073% | 1.776% | 2.555% | 1.711% | 4.072% | 3.779% | 5.566% | 5.075% | 1.667% |  | 3.311% | 1.536% | 3.314% | 3.203%
| PDM3 |  | -1.23% | 27.76% | 1.771% |  | 1.771% | 0.8574% | -26.02% | -8.667% | 9.86% | 3.075% | 9.208% | 6.99% | 10.49% | 6.99% | -0.1472% | 16.12% | 0.3355% |  | -8.537% | -19.65% | -11.73% | 2.535%
| Z-Score (Absolute) |  | 0.16 | 2.01 | NA | 7.1 | NA | 0 | 2.01 | NA | 0.67 | 0.17 | 0.63 | 0.46 | 0.72 | NA | 0.08 | NA | 0.04 | 1.69 | 0.7 | 1.54 | 0.94 | 0.13
Results from laboratories 1.06, 1.08, 1.19 and 1.21 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 235b (Certified Value 0.34 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.31 | 0.32922 | 0.4 | < 2 | 0.41 | 0.33 | 0.35 | 0.3 | 0.31 | 0.35 | 0.32 | 0.36 | 0.3 | 0.35186 | 2 | 0.36 | 0.33 | 38.41 | < 5 | 0.31 | 0.5586 | 0.262 |  |  | 1
| 1 | 2 | 0.34 | 0.34499 | 0.37 | < 2 | 0.43 | 0.31 | 0.34 | 0.3 | 0.37 | 0.34 | 0.31 | 0.42 | 0.33 | 0.31877 | 4 | 0.361 | 0.35 | 43.11 | < 5 | 0.3 | 0.489 | 1.2509 |  |  | 29
| 1 | 3 | 0.32 | 0.31867 | 0.35 | < 2 | 0.44 | 0.31 | 0.35 | 0.3 | 0.32 | 0.35 | 0.31 | 0.36 | 0.3 | 0.35708 | 4 | 0.393 | 0.32 | 10.42 | < 5 | 0.33 | 0.549 | 1.0885 |  |  | 16
| 1 | 4 | 0.32 | 0.32416 | 0.41 | < 2 | 0.44 | 0.33 | 0.36 | 0.3 | 0.31 | 0.33 | 0.31 | 0.36 | 0.31 | 0.33715 | 5 | 0.367 | 0.35 | 24.47 | < 5 | 0.31 | 0.5587 | 1.875 |  |  | 0.33605
| 1 | 5 | 0.33 | 0.34452 | 0.31 | < 2 | 0.41 | 0.32 | 0.4 | 0.3 | 0.33 | 0.33 | 0.34 | 0.37 | 0.32 | 0.33893 | 3 | 0.407 | 0.32 | 44.95 | < 5 | 0.3 | 0.565 | 1.5642 |  |  | 77
| 1 | 6 | 0.34 | 0.33602 | 0.34 | < 2 | 0.44 | 0.34 | 0.33 | 0.4 | 0.3 | 0.32 | 0.32 | 0.36 | 0.31 | 0.35951 | 4 | 0.435 | 0.34 | 45.44 | < 5 | 0.33 | 0.4946 | 0.836
| Mean |  | 0.32667 | 0.33293 | 0.36333 |  | 0.42833 | 0.32333 | 0.355 | 0.31667 | 0.32333 | 0.33667 | 0.31833 | 0.37167 | 0.31167 | 0.34388 | 3.6667 | 0.38717 | 0.335 | 34.4667 |  | 0.31333 | 0.53582 | 1.1461
| Median |  | 0.325 | 0.33262 | 0.36 |  | 0.435 | 0.325 | 0.35 | 0.3 | 0.315 | 0.335 | 0.315 | 0.36 | 0.31 | 0.3454 | 4 | 0.38 | 0.335 | 40.76 |  | 0.31 | 0.5538 | 1.1697
| Std Dev. |  | 0.012111 | 0.010801 | 0.037771 |  | 0.01472 | 0.012111 | 0.02429 | 0.040825 | 0.025033 | 0.012111 | 0.01169 | 0.024014 | 0.01169 | 0.01538 | 1.0328 | 0.030149 | 0.013784 | 14.1333 |  | 0.013663 | 0.034522 | 0.56537
| Rel.Std.Dev. |  | 3.707% | 3.244% | 10.4% |  | 3.436% | 3.746% | 6.842% | 12.89% | 7.742% | 3.597% | 3.672% | 6.461% | 3.751% | 4.472% | 28.17% | 7.787% | 4.115% | 41.01% |  | 4.36% | 6.443% | 49.33%
| PDM3 |  | -2.793% | -0.9294% | 8.118% |  | 27.46% | -3.785% | 5.638% | -5.769% | -3.785% | 0.1825% | -5.273% | 10.6% | -7.257% | 2.33% | 991.1% | 15.21% | -0.3134% | 10156% |  | -6.761% | 59.44% | 241%
| Z-Score (Absolute) |  | 0.54 | 0.39 | 0.33 | 15.33 | 1.86 | 0.62 | 0.13 | NA | 0.62 | 0.3 | 0.73 | 0.52 | 0.89 | 0.13 | NA | 0.89 | 0.34 | 804.16 | 50.69 | 0.85 | 4.39 | 18.78
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 235b (Certified Value 0.335 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.319 | 0.35163 | 0.4 | 0.38833 | 0.37 | 0.35 | 0.34 | 0.35 | 0.3 | 0.34 | 0.34 | 0.33 | 0.33 | 0.36217 | 0.314 | 0.31 | 0.32651 | 0.31 | 0.35 | 0.2789 |  |  |  |  | 1
| 1 | 2 | 0.317 | 0.3504 | 0.42 | 0.38933 | 0.36 | 0.33 | 0.35 | 0.36 | 0.31 | 0.34 | 0.34 | 0.33 | 0.34 | 0.36184 | 0.307 | 0.34 | 0.31942 | 0.31 | 0.35 | 0.2712 |  |  |  |  | 0
| 1 | 3 | 0.327 | 0.34987 | 0.42 | 0.393 | 0.36 | 0.34 | 0.34 | 0.53 | 0.31 | 0.34 | 0.35 | 0.33 | 0.34 | 0.35992 | 0.322 | 0.32 | 0.32375 | 0.31 | 0.33 | 0.2704 |  |  |  |  | 16
| 1 | 4 | 0.323 | 0.34568 | 0.42 | 0.389 | 0.36 | 0.34 | 0.36 | 0.35 | 0.31 | 0.33 | 0.35 | 0.33 | 0.34 | 0.36012 | 0.314 | 0.32 | 0.32153 | 0.31 | 0.33 | 0.2685 |  |  |  |  | 0.33549
| 1 | 5 | 0.324 | 0.34322 | 0.41 | 0.389 | 0.36 | 0.35 | 0.37 | 0.37 | 0.31 | 0.34 | 0.34 | 0.33 | 0.33 | 0.35878 | 0.322 | 0.3 | 0.33176 | 0.31 | 0.33 | 0.276 |  |  |  |  | 78
| 1 | 6 | 0.325 | 0.34417 | 0.42 | 0.388 | 0.36 | 0.35 | 0.35 | 0.36 | 0.31 | 0.34 | 0.35 | 0.33 | 0.35 | 0.35767 | 0.314 | 0.31 | 0.32983 | 0.31 | 0.34 | 0.2788
| Mean |  | 0.3225 | 0.3475 | 0.415 | 0.38944 | 0.36167 | 0.34333 | 0.35167 | 0.38667 | 0.30833 | 0.33833 | 0.345 | 0.33 | 0.33833 | 0.36009 | 0.3155 | 0.31667 | 0.32547 | 0.31 | 0.33833 | 0.27397
| Median |  | 0.3235 | 0.34777 | 0.42 | 0.389 | 0.36 | 0.345 | 0.35 | 0.36 | 0.31 | 0.34 | 0.345 | 0.33 | 0.34 | 0.36002 | 0.314 | 0.315 | 0.32513 | 0.31 | 0.335 | 0.2736
| Std Dev. |  | 0.0037815 | 0.0035698 | 0.0083666 | 0.0018094 | 0.0040825 | 0.008165 | 0.01169 | 0.070616 | 0.0040825 | 0.0040825 | 0.0054772 | 0 | 0.0075277 | 0.0017298 | 0.0057184 | 0.013663 | 0.0047904 | 0 | 0.0098319 | 0.0045187
| Rel.Std.Dev. |  | 1.173% | 1.027% | 2.016% | 0.4646% | 1.129% | 2.378% | 3.324% | 18.26% | 1.324% | 1.207% | 1.588% | 0% | 2.225% | 0.4804% | 1.812% | 4.315% | 1.472% | 0% | 2.906% | 1.649%
| PDM3 |  | -3.872% | 3.579% | 23.7% | 16.08% | 7.803% | 2.338% | 4.822% | 15.25% | -8.094% | 0.848% | 2.835% | -1.636% | 0.848% | 7.332% | -5.958% | -5.61% | -2.988% | -7.597% | 0.848% | -18.34%
| Z-Score (Absolute) |  | 0.57 | 0.33 | 2.76 | 1.84 | 0.84 | 0.18 | 0.48 | 1.74 | 1.08 | 0 | 0.24 | 0.3 | 0 | 0.78 | 0.82 | 0.78 | 0.46 | 1.02 | 0 | 2.31
| Analytical results for Cd in OREAS 235b (Certified Value 0.046 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.037 | 1.18 | 0.04 | 0.7134 | 0.06 | 0.08 | 0.07 | < 0.1 | 0.04 | 0.03 | 0.03 | 0.04 | 0.04 | 0.051003 | < 1 | < 0.1 | 0.05 | 13.21 | 6.5 | 0.03 | 0.94 | 0.3328 |  |  | 1
| 1 | 2 | 0.033 | 1.31 | 0.03 | 0.7 | 0.05 | 0.07 | 0.07 | 0.2 | 0.03 | 0.04 | 0.03 | 0.03 | 0.05 | 0.05171 | < 1 | < 0.1 | 0.04 | 12.74 | 6.5 | 0.04 | 0.9 | 0.348 |  |  | 5
| 1 | 3 | 0.035 | 1.17 | 0.03 | 0.7298 | 0.08 | 0.07 | 0.06 | < 0.1 | 0.04 | 0.04 | 0.03 | 0.04 | 0.05 | 0.050481 | < 1 | < 0.1 | 0.04 | 13.26 | 6.8 | 0.04 | 0.91 | 0.3404 |  |  | 16
| 1 | 4 | 0.037 | 1.25 | 0.03 | 0.7155 | 0.06 | 0.08 | 0.07 | < 0.1 | 0.04 | 0.04 | 0.03 | 0.04 | 0.04 | 0.051455 | < 1 | < 0.1 | 0.04 | 12.89 | 6.3 | 0.04 | 0.85 | 0.3506 |  |  | 0.045903
| 1 | 5 | 0.038 | 1.2 | 0.04 | 0.72015 | 0.06 | 0.08 | 0.07 | < 0.1 | 0.03 | 0.03 | 0.03 | 0.05 | 0.04 | 0.050734 | < 1 | < 0.1 | 0.04 | 11.53 | 6.7 | 0.04 | 0.97 | 0.338 |  |  | 79
| 1 | 6 | 0.034 | 1.24 | 0.04 | 0.74 | 0.08 | 0.08 | 0.06 | < 0.1 | 0.03 | 0.04 | 0.03 | 0.04 | 0.05 | 0.04872 | < 1 | < 0.1 | 0.04 | 11.83 | 6.5 | 0.04 | 0.89 | 0.3431
| Mean |  | 0.035667 | 1.225 | 0.035 | 0.71981 | 0.065 | 0.076667 | 0.066667 | 0.2 | 0.035 | 0.036667 | 0.03 | 0.04 | 0.045 | 0.050684 |  |  | 0.041667 | 12.5767 | 6.55 | 0.038333 | 0.91 | 0.34215
| Median |  | 0.036 | 1.22 | 0.035 | 0.71783 | 0.06 | 0.08 | 0.07 | 0.2 | 0.035 | 0.04 | 0.03 | 0.04 | 0.045 | 0.050869 |  |  | 0.04 | 12.815 | 6.5 | 0.04 | 0.905 | 0.34175
| Std Dev. |  | 0.0019664 | 0.05244 | 0.0054772 | 0.013854 | 0.012247 | 0.005164 | 0.005164 |  | 0.0054772 | 0.005164 | 0 | 0.0063246 | 0.0054772 | 0.0010633 |  |  | 0.0040825 | 0.72745 | 0.17607 | 0.0040825 | 0.041473 | 0.0065458
| Rel.Std.Dev. |  | 5.513% | 4.281% | 15.65% | 1.925% | 18.84% | 6.736% | 7.746% |  | 15.65% | 14.08% | 0% | 15.81% | 12.17% | 2.098% |  |  | 9.798% | 5.784% | 2.688% | 10.65% | 4.557% | 1.913%
| PDM3 |  | -22.3% | 2569% | -23.75% | 1468% | 41.6% | 67.02% | 45.23% | 335.7% | -23.75% | -20.12% | -34.65% | -12.86% | -1.968% | 10.41% |  |  | -9.23% | 27298% | 14169% | -16.49% | 1882% | 645.4%
| Z-Score (Absolute) |  | 0.65 | 51.63 | 0.67 | 29.42 | 0.64 | 1.16 | 0.72 | NA | 0.67 | 0.6 | 0.89 | 0.45 | 0.23 | 0.02 | 19.76 | 0.02 | 0.38 | 550.56 | NA | 0.53 | 37.78 | 12.83
Results from laboratories 1.11 and 1.25 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 235b (Certified Value 56 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 0
| 1 | 1 | 59.838 | 86.0767 | 50.3 | 71.2 | 47.55 | 74.9 | 58 | 53.2 | 53.4 | 59.5 | 58.3 | 56.4731 | 60.14 | 58.73 | 55.79 | 41.1923 | 0.2009 | [7 empty] | 1
| 1 | 2 | 58.531 | 86.0873 | 50.9 | 70.9 | 41.14 | 74.4 | 59 | 56.2 | 54.7 | 58.4 | 59.3 | 56.0331 | 57.59 | 60.83 | 54.68 | 39.1472 | 0.246 | [7 empty] | 31
| 1 | 3 | 59.549 | 85.472 | 50.7 | 70.19 | 42.45 | 75.8 | 57 | 54.9 | 52.8 | 58.5 | 61.3 | 54.7422 | 66.32 | 62.64 | 55.96 | 47.8405 | 0.2473 | [7 empty] | 16
| 1 | 4 | 60.01 | 86.6964 | 49.8 | 69.67 | 48.63 | 75.7 | 56 | 57.4 | 53.2 | 57.6 | 60.4 | 57.118 | 60.21 | 62 | 56.34 | 46.9901 | 0.2381 | [7 empty] | 56.4326
| 1 | 5 | 56.516 | 86.2715 | 52 | 70.1 | 53.42 | 76.8 | 56 | 56.1 | 54.7 | 61.3 | 60.1 | 53.6642 | 57.25 | 63.15 | 54.94 | 46.7569 | 0.2409 | [7 empty] | 80
| 1 | 6 | 58.228 | 86.017 | 49.9 | 71.26 | 52.14 | 77.8 | 57 | 54.2 | 54.4 | 58.8 | 60.8 | 57.8458 | 63.93 | 60.42 | 56.46 | 40.4705 | 0.2494
| Mean |  | 58.7787 | 86.1035 | 50.6 | 70.5533 | 47.555 | 75.9 | 57.1667 | 55.3333 | 53.8667 | 59.0167 | 60.0333 | 55.9794 | 60.9067 | 61.295 | 55.695 | 43.7329 | 0.2371
| Median |  | 59.04 | 86.082 | 50.5 | 70.545 | 48.09 | 75.75 | 57 | 55.5 | 53.9 | 58.65 | 60.25 | 56.2531 | 60.175 | 61.415 | 55.875 | 43.9746 | 0.24345
| Std Dev. |  | 1.3203 | 0.39673 | 0.80994 | 0.65659 | 4.9757 | 1.2426 | 1.169 | 1.5253 | 0.83347 | 1.2766 | 1.0838 | 1.5439 | 3.5746 | 1.6315 | 0.73222 | 3.8666 | 0.018223
| Rel.Std.Dev. |  | 2.246% | 0.4608% | 1.601% | 0.9306% | 10.46% | 1.637% | 2.045% | 2.757% | 1.547% | 2.163% | 1.805% | 2.758% | 5.869% | 2.662% | 1.315% | 8.841% | 7.686%
| PDM3 |  | 4.157% | 52.58% | -10.34% | 25.02% | -15.73% | 34.5% | 1.301% | -1.948% | -4.547% | 4.579% | 6.381% | -0.803% | 7.928% | 8.616% | -1.307% | -22.5% | -99.58%
| Z-Score (Absolute) |  | 0.29 | 5.22 | 1.18 | 2.41 | 1.73 | 3.38 | 0 | 0.33 | 0.59 | 0.33 | 0.52 | 0.21 | 0.67 | 0.74 | 0.27 | 2.42 | 10.26
| Analytical results for Co in OREAS 235b (Certified Value 16.9 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 16.48 | 18.378 | 17.2 | 18.6635 | 16.4 | 16.6 | 17.3 | 15 | 16.6 | 16.9 | 16.8 | 15.9 | 17.1 | 16.6785 | 15 | 15.21 | 17 | 19.74 | 21.4 | 16.1 | 18.905 | 12.2631 | 16.4091 |  | 1
| 1 | 2 | 16.32 | 18.126 | 17 | 18.5188 | 16.5 | 16.2 | 16.8 | 14.9 | 17.1 | 16.6 | 16.1 | 16.4 | 17.2 | 16.6638 | 14 | 13.76 | 16.6 | 16.5 | 20.7 | 16.8 | 17.955 | 13.3688 | 16.4223 |  | 32
| 1 | 3 | 16.12 | 17.757 | 16.6 | 18.5529 | 16.3 | 16.6 | 17.4 | 14.7 | 17.1 | 16.7 | 16.1 | 16.4 | 17.6 | 17.493 | 14 | 23.35 | 17.6 | 20.68 | 21 | 16.1 | 17.765 | 14.0144 | 16.4999 |  | 16
| 1 | 4 | 16.18 | 18.108 | 17.2 | 18.4697 | 16.4 | 16.5 | 17.3 | 14.5 | 17.3 | 16.1 | 16.7 | 16.4 | 17.6 | 17.2201 | 15 | 17.99 | 16.6 | 21.13 | 21.6 | 15.9 | 17.955 | 14.1079 | 16.8468 |  | 16.8668
| 1 | 5 | 16.23 | 17.865 | 16.6 | 18.4743 | 16.3 | 16.4 | 17 | 14.8 | 17.2 | 16.8 | 17.2 | 16.4 | 16.4 | 17.1126 | 15 | 20.48 | 16.2 | 23.46 | 21.7 | 15.7 | 18.335 | 14.5696 | 16.6606 |  | 81
| 1 | 6 | 16.1 | 17.775 | 17.1 | 18.6721 | 16.5 | 16.4 | 17.8 | 14.7 | 17.2 | 16.8 | 16.5 | 16.1 | 16.8 | 17.122 | 15 | 14.27 | 16.5 | 17.08 | 21.4 | 15.6 | 18.525 | 14.1107 | 16.5528
| Mean |  | 16.2383 | 18.0015 | 16.95 | 18.5586 | 16.4 | 16.45 | 17.2667 | 14.7667 | 17.0833 | 16.65 | 16.5667 | 16.2667 | 17.1167 | 17.0483 | 14.6667 | 17.51 | 16.75 | 19.765 | 21.3 | 16.0333 | 18.24 | 13.7391 | 16.5652
| Median |  | 16.205 | 17.9865 | 17.05 | 18.5359 | 16.4 | 16.45 | 17.3 | 14.75 | 17.15 | 16.75 | 16.6 | 16.4 | 17.15 | 17.1173 | 15 | 16.6 | 16.6 | 20.21 | 21.4 | 16 | 18.145 | 14.0611 | 16.5263
| Std Dev. |  | 0.14261 | 0.24426 | 0.28107 | 0.090014 | 0.089443 | 0.15166 | 0.34448 | 0.17512 | 0.24833 | 0.2881 | 0.4274 | 0.21602 | 0.46655 | 0.32291 | 0.5164 | 3.8212 | 0.48888 | 2.6162 | 0.37947 | 0.4274 | 0.42908 | 0.81918 | 0.16592
| Rel.Std.Dev. |  | 0.8782% | 1.357% | 1.658% | 0.485% | 0.5454% | 0.9219% | 1.995% | 1.186% | 1.454% | 1.73% | 2.58% | 1.328% | 2.726% | 1.894% | 3.521% | 21.82% | 2.919% | 13.24% | 1.782% | 2.666% | 2.352% | 5.962% | 1.002%
| PDM3 |  | -3.726% | 6.727% | 0.4931% | 10.03% | -2.768% | -2.471% | 2.371% | -12.45% | 1.284% | -1.285% | -1.78% | -3.558% | 1.481% | 1.076% | -13.04% | 3.813% | -0.6926% | 17.18% | 26.28% | -4.942% | 8.141% | -18.54% | -1.788%
| Z-Score (Absolute) |  | 0.8 | 1.5 | 0.13 | 2.23 | 0.59 | 0.52 | 0.54 | 2.72 | 0.3 | 0.26 | 0.37 | 0.76 | 0.35 | 0.26 | NA | 0.86 | 0.13 | 3.8 | 5.81 | 1.07 | 1.81 | 4.06 | 0.37
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 235b (Certified Value 120 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 128.9 | 117.8855 | 107 | 124.9867 | 112 | 127 | 125 | 122 | 116 | 128 | 121 | 117 | 118.8 | 113.8767 | 121 | 126.1 | 115 | 119.06 | 124 | 116.5 | 82.76 | 123.1996 |  |  | 1
| 1 | 2 | 123.1 | 117.591 | 106 | 124.9033 | 114 | 122 | 124 | 126 | 118 | 126 | 121 | 115 | 121.3 | 114.1688 | 118 | 126.4 | 118 | 120.14 | 124 | 112.5 | 80.3746 | 123.9439 |  |  | 33
| 1 | 3 | 126.1 | 114.8505 | 103 | 126.48 | 116 | 124 | 124 | 126 | 118 | 126 | 122 | 116 | 119.5 | 119.1239 | 121 | 131.5 | 116 | 117.5 | 120 | 112.3 | 79.4494 | 123.7114 |  |  | 16
| 1 | 4 | 125.5 | 119.13 | 109 | 125.3 | 113 | 125 | 124 | 121 | 119 | 125 | 122 | 116 | 118.9 | 117.6853 | 120 | 131.5 | 116 | 119.85 | 119 | 112.2 | 77.181 | 121.9216 |  |  | 120.1468
| 1 | 5 | 124.3 | 115.2825 | 104 | 124.7667 | 116 | 124 | 128 | 124 | 119 | 125 | 122 | 117 | 116.9 | 122.6502 | 122 | 132.1 | 111 | 117.16 | 120 | 114.5 | 89.6644 | 125.2455 |  |  | 82
| 1 | 6 | 123.9 | 116.2135 | 109 | 126.45 | 115 | 126 | 124 | 121 | 117 | 125 | 121 | 117 | 121.6 | 122.0427 | 120 | 129.9 | 113 | 119.42 | 122 | 116.3 | 87.1461 | 123.1612
| Mean |  | 125.3 | 116.8255 | 106.3333 | 125.4811 | 114.3333 | 124.6667 | 124.8333 | 123.3333 | 117.8333 | 125.8333 | 121.5 | 116.3333 | 119.5 | 118.258 | 120.3333 | 129.5833 | 114.8333 | 118.855 | 121.5 | 114.05 | 82.7626 | 123.5305
| Median |  | 124.9 | 116.9022 | 106.5 | 125.1433 | 114.5 | 124.5 | 124 | 123 | 118 | 125.5 | 121.5 | 116.5 | 119.2 | 118.4046 | 120.5 | 130.7 | 115.5 | 119.24 | 121 | 113.5 | 81.5673 | 123.4555
| Std Dev. |  | 2.0707 | 1.6539 | 2.5033 | 0.78206 | 1.633 | 1.7512 | 1.6021 | 2.3381 | 1.169 | 1.169 | 0.54772 | 0.8165 | 1.747 | 3.758 | 1.3663 | 2.6851 | 2.4833 | 1.2419 | 2.1679 | 2.0077 | 4.7892 | 1.0936
| Rel.Std.Dev. |  | 1.653% | 1.416% | 2.354% | 0.6232% | 1.428% | 1.405% | 1.283% | 1.896% | 0.9921% | 0.929% | 0.4508% | 0.7019% | 1.462% | 3.178% | 1.135% | 2.072% | 2.163% | 1.045% | 1.784% | 1.76% | 5.787% | 0.8852%
| PDM3 |  | 4.289% | -2.764% | -11.5% | 4.44% | -4.839% | 3.762% | 3.901% | 2.652% | -1.926% | 4.733% | 1.126% | -3.174% | -0.5384% | -1.572% | 0.1552% | 7.854% | -4.423% | -1.075% | 1.126% | -5.074% | -31.12% | 2.816%
| Z-Score (Absolute) |  | 0.87 | 0.5 | 2.19 | 0.9 | 0.9 | 0.77 | 0.79 | 0.55 | 0.34 | 0.95 | 0.26 | 0.58 | 0.07 | 0.27 | 0.07 | 1.56 | 0.82 | 0.17 | 0.26 | 0.95 | 5.99 | 0.58
| Analytical results for Cs in OREAS 235b (Certified Value 7.42 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.20 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 7.32 | 5.8898 | 7.41 | 7.46 | 7.36 | 8.55 | 7.1 | 7.47 | 7.46 | 7.49 | 7.56 | 7.142 | 7.8 | 7.57 | 7.5351 | [9 empty] | 1
| 1 | 2 | 7.18 | 5.8765 | 7.21 | 7.65 | 7.19 | 8.43 | 7.1 | 7.69 | 7.64 | 7.4 | 7.69 | 6.852 | 7.04 | 7.57 | 6.9539 | [9 empty] | 34
| 1 | 3 | 7.21 | 5.8364 | 7.18 | 7.44 | 7.13 | 8.63 | 7.1 | 7.47 | 7.63 | 7.38 | 7.6 | 7.79 | 7.73 | 7.92 | 8.43 | [9 empty] | 16
| 1 | 4 | 7.38 | 5.7328 | 7.63 | 7.51 | 7.24 | 8.63 | 7 | 7.7 | 7.41 | 7.42 | 7.7 | 7.614 | 7.4 | 7.43 | 8.4041 | [9 empty] | 7.4183
| 1 | 5 | 7.35 | 6.0315 | 7.2 | 7.44 | 7.31 | 8.66 | 7 | 7.6 | 7.69 | 7.78 | 7.86 | 7.812 | 6.97 | 7.32 | 8.3804 | [9 empty] | 83
| 1 | 6 | 7.19 | 5.9815 | 6.94 | 7.68 | 7.5 | 8.82 | 7.1 | 7.53 | 7.46 | 7.26 | 7.6 | 6.989 | 6.98 | 7.66 | 7.4142
| Mean |  | 7.2717 | 5.8914 | 7.2617 | 7.53 | 7.2883 | 8.62 | 7.0667 | 7.5767 | 7.5483 | 7.455 | 7.6683 | 7.3665 | 7.32 | 7.5783 | 7.8529
| Median |  | 7.265 | 5.8831 | 7.205 | 7.485 | 7.275 | 8.63 | 7.1 | 7.565 | 7.545 | 7.41 | 7.645 | 7.378 | 7.22 | 7.57 | 7.9577
| Std Dev. |  | 0.088412 | 0.10599 | 0.23422 | 0.10807 | 0.13227 | 0.12869 | 0.05164 | 0.10347 | 0.11822 | 0.17593 | 0.10889 | 0.42348 | 0.37984 | 0.20605 | 0.6351
| Rel.Std.Dev. |  | 1.216% | 1.799% | 3.225% | 1.435% | 1.815% | 1.493% | 0.7308% | 1.366% | 1.566% | 2.36% | 1.42% | 5.749% | 5.189% | 2.719% | 8.087%
| PDM3 |  | -1.976% | -20.58% | -2.111% | 1.506% | -1.751% | 16.2% | -4.74% | 2.135% | 1.753% | 0.4952% | 3.371% | -0.6978% | -1.325% | 2.158% | 5.86%
| Z-Score (Absolute) |  | 0.74 | 6.33 | 0.78 | 0.3 | 0.67 | 4.71 | 1.57 | 0.49 | 0.38 | 0 | 0.86 | 0.36 | 0.55 | 0.5 | 1.61
| Analytical results for Cu in OREAS 235b (Certified Value 28.4 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 29.66 | 26.312 | 28 | 26.7267 | 30 | 26 | 26 | 26.5 | 28.7 | 28.2 | 29.3 | 29.4 | 28.6 | 30.1396 | 31 | 28.8 | 27.1 | 27.13 | 32.3 | 29.1 | 30.5425 | 27 |  | 28.5042 | 1
| 1 | 2 | 28.06 | 25.898 | 28 | 26.8333 | 29 | 26 | 28 | 27.7 | 28.9 | 29.2 | 28 | 29.6 | 28.48 | 31.0035 | 30 | 27.29 | 29.2 | 27.42 | 31.9 | 30.1 | 28.975 | 27 | 1 | 28.3944 | 35
| 1 | 3 | 29.45 | 24.701 | 26.9 | 27.6233 | 29 | 27 | 29 | 26 | 28.5 | 27.5 | 28.6 | 30.1 | 28.32 | 31.3713 | 30 | 35.63 | 27.1 | 26.24 | 31.6 | 29.5 | 28.8515 | 28 | 1 | 28.6791 | 16
| 1 | 4 | 28.2 | 26.114 | 28.3 | 26.62 | 29 | 25 | 27 | 26.1 | 28.8 | 27.9 | 28.8 | 30.3 | 28.72 | 30.3253 | 30 | 31.16 | 27.3 | 27.6 | 31.6 | 28.1 | 28.5855 | 28 | 1 | 29.2268 | 28.3834
| 1 | 5 | 29.15 | 25.133 | 26.9 | 26.975 | 29 | 26 | 28 | 27.1 | 29 | 29 | 29.8 | 29.5 | 27.83 | 28.7744 | 31 | 38.7 | 26 | 26.7 | 32.4 | 28.7 | 29.5735 | 28 | 1 | 29.0671 | 84
| 1 | 6 | 28.63 | 25.556 | 28.8 | 27.395 | 30 | 29 | 26 | 26.6 | 28.8 | 28.4 | 28.7 | 29.4 | 28.3 | 30.6091 | 30 | 26.52 | 26.6 | 27.14 | 32.5 | 28.2 | 30.21 | 29 | 1 | 28.6821
| Mean |  | 28.8583 | 25.619 | 27.8167 | 27.0289 | 29.3333 | 26.5 | 27.3333 | 26.6667 | 28.7833 | 28.3667 | 28.8667 | 29.7167 | 28.375 | 30.3705 | 30.3333 | 31.35 | 27.2167 | 27.0383 | 32.05 | 28.95 | 29.4563 | 27.8333 | 1 | 28.759
| Median |  | 28.89 | 25.727 | 28 | 26.9042 | 29 | 26 | 27.5 | 26.55 | 28.8 | 28.3 | 28.75 | 29.55 | 28.4 | 30.4672 | 30 | 29.98 | 27.1 | 27.135 | 32.1 | 28.9 | 29.2742 | 28 | 1 | 28.6806
| Std Dev. |  | 0.66343 | 0.61417 | 0.7679 | 0.39676 | 0.5164 | 1.3784 | 1.2111 | 0.64083 | 0.17224 | 0.64704 | 0.6186 | 0.38687 | 0.31188 | 0.90133 | 0.5164 | 4.8722 | 1.0797 | 0.49632 | 0.40373 | 0.77395 | 0.7896 | 0.75277 | 0 | 0.32371
| Rel.Std.Dev. |  | 2.299% | 2.397% | 2.761% | 1.468% | 1.76% | 5.202% | 4.431% | 2.403% | 0.5984% | 2.281% | 2.143% | 1.302% | 1.099% | 2.968% | 1.702% | 15.54% | 3.967% | 1.836% | 1.26% | 2.673% | 2.681% | 2.705% | 0% | 1.126%
| PDM3 |  | 1.673% | -9.739% | -1.997% | -4.772% | 3.347% | -6.636% | -3.7% | -6.048% | 1.409% | -0.05891% | 1.703% | 4.697% | -0.02955% | 7.001% | 6.87% | 10.45% | -4.111% | -4.739% | 12.92% | 1.996% | 3.78% | -1.938% | -96.48% | 1.323%
| Z-Score (Absolute) |  | 0.16 | 1.67 | 0.42 | 0.87 | 0.43 | 1.17 | 0.7 | 1.08 | 0.12 | 0.11 | 0.17 | 0.65 | 0.11 | 1.02 | 1 | 1.57 | 0.76 | 0.86 | 1.97 | 0.22 | 0.5 | 0.42 | 15.6 | 0.11
| Analytical results for Dy in OREAS 235b (Indicative Value 2.14 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 2.2325; 1.06: 2.2; 1.07: 2.19; 1.08: 2.0134 |  |  |  | 1
| 1 | 2 | 1.01: 2.2108; 1.06: 2; 1.07: 2.108; 1.08: 1.8715 |  |  |  | 9
| 1 | 3 | 1.01: 2.1884; 1.06: 2; 1.07: 2.294; 1.08: 2.2203 |  |  |  | 16
| 1 | 4 | 1.01: 2.1938; 1.06: 2.1; 1.07: 2.185; 1.08: 2.2225 |  |  |  | 2.1437
| 1 | 5 | 1.01: 2.2866; 1.06: 2.1; 1.07: 2.143; 1.08: 2.2359 |  |  |  | 15
| 1 | 6 | 1.01: 2.2015; 1.06: 2; 1.07: 2.306; 1.08: 1.9447
| Mean |  | 1.01: 2.2189; 1.06: 2.0667; 1.07: 2.2043; 1.08: 2.0847
| Median |  | 1.01: 2.2062; 1.06: 2.05; 1.07: 2.1875; 1.08: 2.1168
| Std Dev. |  | 1.01: 0.036619; 1.06: 0.08165; 1.07: 0.079997; 1.08: 0.16148
| Rel.Std.Dev. |  | 1.01: 1.65%; 1.06: 3.951%; 1.07: 3.629%; 1.08: 7.746%
| PDM3 |  | 1.01: 3.511%; 1.06: -3.592%; 1.07: 2.83%; 1.08: -2.75%
| Z-Score (Absolute) |  | 1.01: 0.75; 1.06: 0.78; 1.07: 0.6; 1.08: 0.6
| Analytical results for Er in OREAS 235b (Indicative Value 0.98 ppm) | [25 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 1.0876; 1.06: 0.9; 1.07: 1.007; 1.08: 0.913 |  |  |  | 1
| 1 | 2 | 1.01: 1.0437; 1.06: 1; 1.07: 0.976; 1.08: 0.8292 |  |  |  | 10
| 1 | 3 | 1.01: 1.0363; 1.06: 0.9; 1.07: 1.083; 1.08: 0.9975 |  |  |  | 16
| 1 | 4 | 1.01: 0.95737; 1.06: 1; 1.07: 1.009; 1.08: 1.016 |  |  |  | 0.97779
| 1 | 5 | 1.01: 0.99233; 1.06: 0.9; 1.07: 0.989; 1.08: 1.0001 |  |  |  | 16
| 1 | 6 | 1.01: 0.99436; 1.06: 0.9; 1.07: 1.068; 1.08: 0.8674
| Mean |  | 1.01: 1.0186; 1.06: 0.93333; 1.07: 1.022; 1.08: 0.9372
| Median |  | 1.01: 1.0153; 1.06: 0.9; 1.07: 1.008; 1.08: 0.95525
| Std Dev. |  | 1.01: 0.046332; 1.06: 0.05164; 1.07: 0.043442; 1.08: 0.078643
| Rel.Std.Dev. |  | 1.01: 4.549%; 1.06: 5.533%; 1.07: 4.251%; 1.08: 8.391%
| PDM3 |  | 1.01: 4.175%; 1.06: -4.546%; 1.07: 4.522%; 1.08: -4.151%
| Z-Score (Absolute) |  | 1.01: 0.65; 1.06: 0.71; 1.07: 0.7; 1.08: 0.65
| Analytical results for Eu in OREAS 235b (Indicative Value 0.62 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.76733; 1.06: 0.6; 1.07: 0.639; 1.08: 0.5266 |  |  |  | 1
| 1 | 2 | 1.01: 0.76352; 1.06: 0.5; 1.07: 0.614; 1.08: 0.5124 |  |  |  | 11
| 1 | 3 | 1.01: 0.72186; 1.06: 0.6; 1.07: 0.698; 1.08: 0.6041 |  |  |  | 16
| 1 | 4 | 1.01: 0.73657; 1.06: 0.6; 1.07: 0.641; 1.08: 0.5926 |  |  |  | 0.62291
| 1 | 5 | 1.01: 0.72004; 1.06: 0.5; 1.07: 0.615; 1.08: 0.5889 |  |  |  | 17
| 1 | 6 | 1.01: 0.72546; 1.06: 0.5; 1.07: 0.676; 1.08: 0.5075
| Mean |  | 1.01: 0.73913; 1.06: 0.55; 1.07: 0.64717; 1.08: 0.55535
| Median |  | 1.01: 0.73102; 1.06: 0.55; 1.07: 0.64; 1.08: 0.55775
| Std Dev. |  | 1.01: 0.021193; 1.06: 0.054772; 1.07: 0.03363; 1.08: 0.044386
| Rel.Std.Dev. |  | 1.01: 2.867%; 1.06: 9.959%; 1.07: 5.196%; 1.08: 7.992%
| PDM3 |  | 1.01: 18.66%; 1.06: -11.71%; 1.07: 3.894%; 1.08: -10.85%
| Z-Score (Absolute) |  | 1.01: 1.91; 1.06: 0.71; 1.07: 0.64; 1.08: 0.64
| Analytical results for Fe in OREAS 235b (Certified Value 3.67 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.609 | 3.7559 | 3.59 | 3.7833 | 3.75 | 3.71 | 4.06 | 3.73 | 3.52 | 3.8 | 3.64 | 3.52 | 3.72 | 3.7513 | 3.672 | 3.52 | 3.6284 | 3.74 | 3.82 | 3.5496 | 3.86 |  |  |  | 1
| 1 | 2 | 3.559 | 3.7426 | 3.59 | 3.76 | 3.77 | 3.64 | 4.02 | 3.73 | 3.58 | 3.76 | 3.63 | 3.47 | 3.82 | 3.8579 | 3.609 | 3.71 | 3.5482 | 3.73 | 3.8 | 3.4457 | 4.2 |  |  |  | 0
| 1 | 3 | 3.643 | 3.7059 | 3.51 | 3.8033 | 3.68 | 3.63 | 4.06 | 3.71 | 3.55 | 3.76 | 3.67 | 3.48 | 3.75 | 3.8361 | 3.714 | 3.57 | 3.596 | 3.74 | 3.7 | 3.4363 | 4.1657 |  |  |  | 16
| 1 | 4 | 3.627 | 3.7817 | 3.64 | 3.78 | 3.68 | 3.75 | 4.06 | 3.69 | 3.61 | 3.67 | 3.67 | 3.51 | 3.74 | 3.8451 | 3.616 | 3.57 | 3.5502 | 3.72 | 3.65 | 3.4035 | 4.1812 |  |  |  | 3.6749
| 1 | 5 | 3.591 | 3.6973 | 3.54 | 3.7667 | 3.69 | 3.73 | 4.16 | 3.7 | 3.58 | 3.7 | 3.66 | 3.57 | 3.65 | 3.8435 | 3.707 | 3.41 | 3.6423 | 3.74 | 3.69 | 3.4954 | 4.1616 |  |  |  | 85
| 1 | 6 | 3.546 | 3.7062 | 3.62 | 3.7833 | 3.8 | 3.74 | 4.08 | 3.74 | 3.56 | 3.71 | 3.67 | 3.53 | 3.83 | 3.7607 | 3.665 | 3.44 | 3.6336 | 3.76 | 3.77 | 3.55 | 4.1796
| Mean |  | 3.5958 | 3.7316 | 3.5817 | 3.7794 | 3.7283 | 3.7 | 4.0733 | 3.7167 | 3.5667 | 3.7333 | 3.6567 | 3.5133 | 3.7517 | 3.8158 | 3.6638 | 3.5367 | 3.5998 | 3.7383 | 3.7383 | 3.4801 | 4.1247
| Median |  | 3.6 | 3.7244 | 3.59 | 3.7817 | 3.72 | 3.72 | 4.06 | 3.72 | 3.57 | 3.735 | 3.665 | 3.515 | 3.745 | 3.8398 | 3.6685 | 3.545 | 3.6122 | 3.74 | 3.735 | 3.4705 | 4.1726
| Std Dev. |  | 0.038034 | 0.03377 | 0.048751 | 0.015117 | 0.051929 | 0.052154 | 0.046762 | 0.019664 | 0.030768 | 0.048028 | 0.017512 | 0.036148 | 0.066758 | 0.046893 | 0.044143 | 0.10764 | 0.042178 | 0.013292 | 0.067946 | 0.061514 | 0.13037
| Rel.Std.Dev. |  | 1.058% | 0.905% | 1.361% | 0.4% | 1.393% | 1.41% | 1.148% | 0.5291% | 0.8627% | 1.286% | 0.4789% | 1.029% | 1.779% | 1.229% | 1.205% | 3.044% | 1.172% | 0.3555% | 1.818% | 1.768% | 3.161%
| PDM3 |  | -2.151% | 1.543% | -2.537% | 2.845% | 1.454% | 0.6832% | 10.84% | 1.137% | -2.945% | 1.59% | -0.4959% | -4.396% | 2.089% | 3.834% | -0.3009% | -3.761% | -2.044% | 1.726% | 1.726% | -5.301% | 12.24%
| Z-Score (Absolute) |  | 1.3 | 0.16 | 1.45 | 0.67 | 0.13 | 0.18 | 3.83 | 0 | 1.61 | 0.18 | 0.64 | 2.18 | 0.38 | 1.06 | 0.57 | 1.93 | 1.26 | 0.23 | 0.23 | 2.54 | 4.38
| Analytical results for Ga in OREAS 235b (Certified Value 9.46 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.27 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 9.6 | 8.5263 | 9.09 | 9.92 | 9.9 | 12 | 9.5 | 10 | 8.83 | 9.31 | 9.82 | 9.24 | 10.1 | 9.7479 | 10 | 12.3 | 10.55 | 8.3947 | 9.2749 | [5 empty] | 1
| 1 | 2 | 9.4 | 8.6915 | 9 | 9.87 | 9.8 | 10 | 9.2 | 10 | 9.24 | 9.56 | 9.37 | 9.27 | 9.8 | 9.5029 | 9.7 | 12.2 | 9.93 | 8.0584 | 8.9923 | [5 empty] | 39
| 1 | 3 | 9.7 | 8.5463 | 9.08 | 9.9 | 9.5 | 10 | 9.2 | 10 | 9.22 | 9.05 | 9.57 | 9.34 | 9.9 | 9.8821 | 9.9 | 12.1 | 10.28 | 9.5972 | 9.3828 | [5 empty] | 16
| 1 | 4 | 9.7 | 8.5869 | 9.55 | 10.0267 | 9.6 | 11 | 9.3 | 10 | 9.3 | 9.02 | 9.71 | 9.4 | 9.9 | 9.6779 | 10 | 12 | 9.97 | 9.6448 | 9.1243 | [5 empty] | 9.4568
| 1 | 5 | 9.6 | 8.3252 | 9.1 | 9.9653 | 9.7 | 12 | 9.3 | 10 | 9.22 | 9.41 | 9.97 | 9.33 | 8.9 | 9.4692 | 9.3 | 12.5 | 9.72 | 9.6475 | 8.9594 | [5 empty] | 86
| 1 | 6 | 9.6 | 8.5023 | 8.95 | 10.1627 | 10 | 11 | 9.7 | 10 | 9.11 | 9.34 | 9.63 | 9.32 | 10 | 9.6344 | 9.8 | 12.3 | 10.25 | 8.3667 | 8.9029
| Mean |  | 9.6 | 8.5298 | 9.1283 | 9.9741 | 9.75 | 11 | 9.3667 | 10 | 9.1533 | 9.2817 | 9.6783 | 9.3167 | 9.7667 | 9.6524 | 9.7833 | 12.2333 | 10.1167 | 8.9515 | 9.1061
| Median |  | 9.6 | 8.5363 | 9.085 | 9.9427 | 9.75 | 11 | 9.3 | 10 | 9.22 | 9.325 | 9.67 | 9.325 | 9.9 | 9.6562 | 9.85 | 12.25 | 10.11 | 8.996 | 9.0583
| Std Dev. |  | 0.10954 | 0.12025 | 0.21479 | 0.10738 | 0.18708 | 0.89443 | 0.19664 | 0 | 0.1699 | 0.2099 | 0.2075 | 0.056095 | 0.43665 | 0.15413 | 0.26394 | 0.17512 | 0.29864 | 0.75255 | 0.19034
| Rel.Std.Dev. |  | 1.141% | 1.41% | 2.353% | 1.077% | 1.919% | 8.131% | 2.099% | 0% | 1.856% | 2.261% | 2.144% | 0.6021% | 4.471% | 1.597% | 2.698% | 1.431% | 2.952% | 8.407% | 2.09%
| PDM3 |  | 1.514% | -9.803% | -3.474% | 5.47% | 3.1% | 16.32% | -0.9534% | 5.744% | -3.209% | -1.852% | 2.342% | -1.482% | 3.276% | 2.068% | 3.453% | 29.36% | 6.977% | -5.343% | -3.709%
| Z-Score (Absolute) |  | 0.05 | 2.14 | 0.97 | 0.68 | 0.24 | 2.68 | 0.51 | NA | 0.92 | 0.67 | 0.1 | 0.6 | 0.27 | 0.05 | 0.31 | 5.08 | 0.96 | 1.31 | 1.01
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 235b (Indicative Value 4.15 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 5.8287; 1.06: 3.2; 1.07: 3.923; 1.08: 3.6159 |  |  |  | 1
| 1 | 2 | 1.01: 5.7051; 1.06: 3.3; 1.07: 3.798; 1.08: 3.381 |  |  |  | 12
| 1 | 3 | 1.01: 5.5489; 1.06: 3.1; 1.07: 4.236; 1.08: 4.0648 |  |  |  | 16
| 1 | 4 | 1.01: 5.5537; 1.06: 3; 1.07: 3.975; 1.08: 4.118 |  |  |  | 4.1458
| 1 | 5 | 1.01: 5.6097; 1.06: 3.1; 1.07: 3.829; 1.08: 4.027 |  |  |  | 18
| 1 | 6 | 1.01: 5.5538; 1.06: 3.3; 1.07: 4.177; 1.08: 3.5536
| Mean |  | 1.01: 5.6333; 1.06: 3.1667; 1.07: 3.9897; 1.08: 3.7934
| Median |  | 1.01: 5.5817; 1.06: 3.15; 1.07: 3.949; 1.08: 3.8215
| Std Dev. |  | 1.01: 0.11285; 1.06: 0.12111; 1.07: 0.18055; 1.08: 0.3139
| Rel.Std.Dev. |  | 1.01: 2.003%; 1.06: 3.824%; 1.07: 4.525%; 1.08: 8.275%
| PDM3 |  | 1.01: 35.88%; 1.06: -23.62%; 1.07: -3.765%; 1.08: -8.5%
| Z-Score (Absolute) |  | 1.01: 2.85; 1.06: 1.19; 1.07: 0.16; 1.08: 0.16
| Analytical results for Ge in OREAS 235b (Certified Value 0.12 ppm) | [25 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.06; 1.07: 1.09; 1.08: 1.11; 1.09: 1.12; 1.10: 1.13; 1.11: 1.14; 1.12: 1.15; 1.13: 1.21; 1.14: 1.27; 1.15: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*OES/MS; 1.09: AR*MS; 1.10: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*OES/MS; 1.15: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.2g; 1.08: 0.25g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.25g; 1.14: 0.5g; 1.15: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.14; 1.08: 1.7; 1.09: 0.19; 1.10: 0.13; 1.11: 0.08; 1.12: 0.12; 1.13: 0.2; 1.14: 0.13; 1.15: 0.0737 |  |  |  | 1
| 1 | 2 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.13; 1.08: 1.7; 1.09: 0.2; 1.10: 0.13; 1.11: 0.08; 1.12: 0.12; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0593 |  |  |  | 6
| 1 | 3 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.13; 1.08: 1.7; 1.09: 0.19; 1.10: 0.12; 1.11: 0.07; 1.12: 0.12; 1.13: 0.2; 1.14: 0.14; 1.15: 0.0837 |  |  |  | 16
| 1 | 4 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.13; 1.08: 1.8; 1.09: 0.2; 1.10: 0.12; 1.11: 0.07; 1.12: 0.13; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0951 |  |  |  | 0.11531
| 1 | 5 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.13; 1.08: 1.7; 1.09: 0.19; 1.10: 0.11; 1.11: 0.08; 1.12: 0.12; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0898 |  |  |  | 87
| 1 | 6 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.13; 1.08: 1.7; 1.09: 0.21; 1.10: 0.11; 1.11: 0.07; 1.12: 0.12; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0676
| Mean |  | 1.07: 0.13167; 1.08: 1.7167; 1.09: 0.19667; 1.10: 0.12; 1.11: 0.075; 1.12: 0.12167; 1.13: 0.2; 1.14: 0.125; 1.15: 0.0782
| Median |  | 1.07: 0.13; 1.08: 1.7; 1.09: 0.195; 1.10: 0.12; 1.11: 0.075; 1.12: 0.12; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0787
| Std Dev. |  | 1.07: 0.0040825; 1.08: 0.040825; 1.09: 0.008165; 1.10: 0.0089443; 1.11: 0.0054772; 1.12: 0.0040825; 1.13: 3.0405e-17; 1.14: 0.0083666; 1.15: 0.013713
| Rel.Std.Dev. |  | 1.07: 3.101%; 1.08: 2.378%; 1.09: 4.152%; 1.10: 7.454%; 1.11: 7.303%; 1.12: 3.355%; 1.13: 1.52e-14%; 1.14: 6.693%; 1.15: 17.54%
| PDM3 |  | 1.07: 14.19%; 1.08: 1389%; 1.09: 70.56%; 1.10: 4.07%; 1.11: -34.96%; 1.12: 5.516%; 1.13: 73.45%; 1.14: 8.407%; 1.15: -32.18%
| Z-Score (Absolute) |  | 1.01: 1.53; 1.06: 1.13; 1.07: 0.19; 1.08: NA; 1.09: 1.24; 1.10: 0; 1.11: 0.73; 1.12: 0.03; 1.13: NA; 1.14: 0.08; 1.15: 0.67
Results from laboratories 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 235b (Certified Value 0.49 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.48 | 0.81322 | 0.2 | 0.65 | 0.52 | 0.69 | 0.5 | 0.46 | 0.44 | 0.48 | 0.51 | 0.53 | 0.32 | 0.2902 | 0.4619 | [9 empty] | 1
| 1 | 2 | 0.5 | 0.78474 | 0.1 | 0.69 | 0.46 | 0.7 | 0.5 | 0.5 | 0.45 | 0.46 | 0.51 | 0.55 | 0.35 | 0.246 | 0.4688 | [9 empty] | 25
| 1 | 3 | 0.5 | 0.78134 | 0.2 | 0.64 | 0.47 | 0.67 | 0.5 | 0.48 | 0.45 | 0.47 | 0.52 | 0.53 | 0.31 | 0.302 | 0.462 | [9 empty] | 16
| 1 | 4 | 0.44 | 0.7917 | 0.2 | 0.62 | 0.59 | 0.72 | 0.5 | 0.49 | 0.45 | 0.45 | 0.52 | 0.55 | 0.32 | 0.2943 | 0.4538 | [9 empty] | 0.48776
| 1 | 5 | 0.47 | 0.81293 | 0.1 | 0.67 | 0.53 | 0.74 | 0.5 | 0.5 | 0.45 | 0.48 | 0.51 | 0.47 | 0.31 | 0.2856 | 0.421 | [9 empty] | 88
| 1 | 6 | 0.43 | 0.78253 | 0.3 | 0.68 | 0.56 | 0.69 | 0.5 | 0.49 | 0.46 | 0.46 | 0.54 | 0.47 | 0.31 | 0.2523 | 0.4689
| Mean |  | 0.47 | 0.79441 | 0.18333 | 0.65833 | 0.52167 | 0.70167 | 0.5 | 0.48667 | 0.45 | 0.46667 | 0.51833 | 0.51667 | 0.32 | 0.2784 | 0.45607
| Median |  | 0.475 | 0.78822 | 0.2 | 0.66 | 0.525 | 0.695 | 0.5 | 0.49 | 0.45 | 0.465 | 0.515 | 0.53 | 0.315 | 0.2879 | 0.46195
| Std Dev. |  | 0.029665 | 0.014896 | 0.075277 | 0.026394 | 0.050365 | 0.024833 | 0 | 0.015055 | 0.0063246 | 0.012111 | 0.01169 | 0.037238 | 0.015492 | 0.023374 | 0.01806
| Rel.Std.Dev. |  | 6.312% | 1.875% | 41.06% | 4.009% | 9.655% | 3.539% | 0% | 3.094% | 1.405% | 2.595% | 2.255% | 7.207% | 4.841% | 8.396% | 3.96%
| PDM3 |  | -3.642% | 62.87% | -62.41% | 34.97% | 6.95% | 43.85% | 2.508% | -0.2251% | -7.742% | -4.325% | 6.267% | 5.925% | -34.39% | -42.92% | -6.499%
| Z-Score (Absolute) |  | 0.32 | 5.93 | 5.84 | 3.31 | 0.67 | 4.14 | 0.26 | 0 | 0.71 | 0.39 | 0.61 | 0.58 | 3.21 | 4.01 | 0.59
| Analytical results for Hg in OREAS 235b (Indicative Value 0.013 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.01 | 0.06 | < 0.05 | < 2 | 0.105 | < 0.01 | 0.02 | 0.01 | < 0.01 | 0.021 | < 1 | < 0.01 | < 5 | < 0.005 | 0.0133 | [9 empty] | 1
| 1 | 2 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.101 | < 0.01 | 0.01 | 0.01 | < 0.01 | 0.017 | < 1 | < 0.01 | < 5 | 0.007 | 0.0168 | [9 empty] | 13
| 1 | 3 | < 0.01 | 0.04 | < 0.05 | < 2 | 0.102 | 0.01 | 0.01 | 0.01 | < 0.01 | 0.018 | < 1 | < 0.01 | < 5 | < 0.005 | 0.0125 | [9 empty] | 16
| 1 | 4 | < 0.01 | 0.05 | < 0.05 | < 2 | 0.098 | 0.01 | 0.02 | 0.01 | < 0.01 | 0.021 | < 1 | < 0.01 | < 5 | 0.006 | 0.0166 | [9 empty] | 0.012911
| 1 | 5 | < 0.01 | 0.03 | < 0.05 | < 2 | 0.107 | 0.01 | 0.02 | 0.01 | < 0.01 | 0.019 | < 1 | < 0.01 | < 5 | 0.006 | 0.0202 | [9 empty] | 19
| 1 | 6 | < 0.01 | 0.06 | < 0.05 | < 2 | 0.106 | 0.01 | 0.02 | 0.01 | < 0.01 | 0.018 | < 1 | < 0.01 | < 5 | 0.01 | 0.0079
| Mean |  | 1.06: 0.048333; 1.09: 0.10317; 1.10: 0.01; 1.11: 0.016667; 1.12: 0.01; 1.14: 0.019; 1.19: 0.00725; 1.21: 0.01455
| Median |  | 1.06: 0.05; 1.09: 0.1035; 1.10: 0.01; 1.11: 0.02; 1.12: 0.01; 1.14: 0.0185; 1.19: 0.0065; 1.21: 0.01495
| Std Dev. |  | 1.06: 0.01169; 1.09: 0.0034303; 1.10: 0; 1.11: 0.005164; 1.12: 0; 1.14: 0.0016733; 1.19: 0.001893; 1.21: 0.0042702
| Rel.Std.Dev. |  | 1.06: 24.19%; 1.09: 3.325%; 1.10: 0%; 1.11: 30.98%; 1.12: 0%; 1.14: 8.807%; 1.19: 26.11%; 1.21: 29.35%
| PDM3 |  | 1.06: 274.4%; 1.09: 699.1%; 1.10: -22.55%; 1.11: 29.09%; 1.12: -22.55%; 1.14: 47.16%; 1.19: -43.85%; 1.21: 12.69%
| Z-Score (Absolute) |  | 0.67 | 1.83 | 0.48 | 56.83 | 5 | 0.48 | 0 | 0.39 | 0.67 | 0.13 | 27.94 | 0.67 | 143.53 | 0.64 | 0.12
| Analytical results for Ho in OREAS 235b (Indicative Value 0.36 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.37583; 1.06: 0.4; 1.07: 0.374; 1.08: 0.3428 |  |  |  | 1
| 1 | 2 | 1.01: 0.39099; 1.06: 0.4; 1.07: 0.362; 1.08: 0.3151 |  |  |  | 14
| 1 | 3 | 1.01: 0.38311; 1.06: 0.3; 1.07: 0.395; 1.08: 0.3769 |  |  |  | 16
| 1 | 4 | 1.01: 0.37707; 1.06: 0.3; 1.07: 0.375; 1.08: 0.3804 |  |  |  | 0.36191
| 1 | 5 | 1.01: 0.38934; 1.06: 0.3; 1.07: 0.367; 1.08: 0.3814 |  |  |  | 20
| 1 | 6 | 1.01: 0.37536; 1.06: 0.3; 1.07: 0.393; 1.08: 0.3315
| Mean |  | 1.01: 0.38195; 1.06: 0.33333; 1.07: 0.37767; 1.08: 0.35468
| Median |  | 1.01: 0.38009; 1.06: 0.3; 1.07: 0.3745; 1.08: 0.35985
| Std Dev. |  | 1.01: 0.0069638; 1.06: 0.05164; 1.07: 0.01353; 1.08: 0.028685
| Rel.Std.Dev. |  | 1.01: 1.823%; 1.06: 15.49%; 1.07: 3.583%; 1.08: 8.088%
| PDM3 |  | 1.01: 5.538%; 1.06: -7.896%; 1.07: 4.354%; 1.08: -1.996%
| Z-Score (Absolute) |  | 1.01: 0.78; 1.06: 1.62; 1.07: 0.57; 1.08: 0.57
| Analytical results for In in OREAS 235b (Certified Value 0.037 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.04 | 0.04 | < 5 | < 0.05 | 0.021 | < 0.05 | 0.04 | 0.035 | 0.034 | 0.034 | 0.035 | 0.04 | 0.036 | 0.0399 | [10 empty] | 1
| 1 | 2 | 0.04 | 0.03 | < 5 | < 0.05 | 0.024 | < 0.05 | 0.04 | 0.039 | 0.036 | 0.036 | 0.037 | 0.04 | 0.034 | 0.0365 | [10 empty] | 27
| 1 | 3 | 0.04 | 0.04 | < 5 | < 0.05 | 0.029 | < 0.05 | 0.04 | 0.036 | 0.035 | 0.036 | 0.038 | 0.04 | 0.036 | 0.0432 | [10 empty] | 16
| 1 | 4 | 0.04 | 0.02 | < 5 | < 0.05 | 0.022 | < 0.05 | 0.04 | 0.037 | 0.037 | 0.038 | 0.038 | 0.04 | 0.034 | 0.0502 | [10 empty] | 0.037185
| 1 | 5 | 0.04 | 0.04 | < 5 | < 0.05 | 0.025 | < 0.05 | 0.04 | 0.038 | 0.036 | 0.037 | 0.038 | 0.03 | 0.033 | 0.0457 | [10 empty] | 89
| 1 | 6 | 0.03 | 0.05 | < 5 | < 0.05 | 0.03 | < 0.05 | 0.04 | 0.04 | 0.035 | 0.035 | 0.036 | 0.04 | 0.034 | 0.0365
| Mean |  | 1.01: 0.038333; 1.06: 0.036667; 1.09: 0.025167; 1.11: 0.04; 1.12: 0.0375; 1.13: 0.0355; 1.14: 0.036; 1.15: 0.037; 1.16: 0.038333; 1.17: 0.0345; 1.19: 0.042
| Median |  | 1.01: 0.04; 1.06: 0.04; 1.09: 0.0245; 1.11: 0.04; 1.12: 0.0375; 1.13: 0.0355; 1.14: 0.036; 1.15: 0.0375; 1.16: 0.04; 1.17: 0.034; 1.19: 0.04155
| Std Dev. |  | 1.01: 0.0040825; 1.06: 0.010328; 1.09: 0.003656; 1.11: 0; 1.12: 0.0018708; 1.13: 0.0010488; 1.14: 0.0014142; 1.15: 0.0012649; 1.16: 0.0040825; 1.17: 0.0012247; 1.19: 0.0054273
| Rel.Std.Dev. |  | 1.01: 10.65%; 1.06: 28.17%; 1.09: 14.53%; 1.11: 0%; 1.12: 4.989%; 1.13: 2.954%; 1.14: 3.928%; 1.15: 3.419%; 1.16: 10.65%; 1.17: 3.55%; 1.19: 12.92%
| PDM3 |  | 1.01: 3.088%; 1.06: -1.394%; 1.09: -32.32%; 1.11: 7.57%; 1.12: 0.8466%; 1.13: -4.532%; 1.14: -3.187%; 1.15: -0.498%; 1.16: 3.088%; 1.17: -7.221%; 1.19: 12.95%
| Z-Score (Absolute) |  | 0.53 | 0.06 | 866.57 | 4.16 | 4.1 | 4.16 | 1.11 | 0.23 | 0.47 | 0.29 | 0.06 | 0.53 | 0.82 | 1.82
| Analytical results for Ir in OREAS 235b (Indicative Value < 0.0016 ppm) | [25 empty] | Indicative
Ir | Round | Replicate | 1.01: Lab |  |  |  | 1
| No | No | 1.01: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS |  |  |  | 3
|  |  | 1.01: 0.5g |  |  |  | 3
| 1 | 1 | 1.01: < 0.0016 |  |  |  | 1
| 1 | 2 | 1.01: < 0.0016 |  |  |  | 4
| 1 | 3 | 1.01: < 0.0018 |  |  |  | 16
| 1 | 4 | 1.01: < 0.0016 |  |  |  | < 0.0016
| 1 | 5 | 1.01: < 0.0019 |  |  |  | 14
| 1 | 6 | 1.01: < 0.001
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for K in OREAS 235b (Certified Value 0.9 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.834 | 1.0635 | 0.79 | 0.9195 | 0.945 | 0.93 | 0.96 | 0.86 | 0.84 | 0.96 | 0.92 | 0.87 | 0.89 | 0.88269 | 0.897 | 0.86 | 0.87426 | 0.9 | 0.95 | 0.9235 | 1.0071 |  |  |  | 1
| 1 | 2 | 0.838 | 1.0471 | 0.79 | 0.9155 | 0.944 | 0.89 | 0.965 | 0.87 | 0.85 | 0.95 | 0.93 | 0.85 | 0.92 | 0.88621 | 0.872 | 0.92 | 0.85056 | 0.9 | 0.94 | 0.9123 | 1.0067 |  |  |  | 0
| 1 | 3 | 0.848 | 1.0109 | 0.79 | 0.921 | 0.932 | 0.88 | 0.96 | 0.86 | 0.86 | 0.94 | 0.93 | 0.86 | 0.9 | 0.87557 | 0.897 | 0.86 | 0.85604 | 0.89 | 0.91 | 0.9123 | 1.0073 |  |  |  | 16
| 1 | 4 | 0.839 | 1.0153 | 0.81 | 0.9105 | 0.946 | 0.89 | 0.988 | 0.85 | 0.87 | 0.94 | 0.94 | 0.87 | 0.89 | 0.89847 | 0.88 | 0.87 | 0.85169 | 0.9 | 0.89 | 0.9019 | 1.0065 |  |  |  | 0.89959
| 1 | 5 | 0.862 | 1.0295 | 0.8 | 0.919 | 0.934 | 0.91 | 1.02 | 0.86 | 0.87 | 0.95 | 0.93 | 0.88 | 0.87 | 0.87246 | 0.905 | 0.82 | 0.87591 | 0.9 | 0.9 | 0.9263 | 1.0066 |  |  |  | 90
| 1 | 6 | 0.843 | 1.0615 | 0.83 | 0.92 | 0.953 | 0.93 | 0.978 | 0.87 | 0.86 | 0.95 | 0.94 | 0.87 | 0.92 | 0.85236 | 0.897 | 0.84 | 0.87414 | 0.9 | 0.93 | 0.9329 | 1.0065
| Mean |  | 0.844 | 1.038 | 0.80167 | 0.91758 | 0.94233 | 0.905 | 0.9785 | 0.86167 | 0.85833 | 0.94833 | 0.93167 | 0.86667 | 0.89833 | 0.87796 | 0.89133 | 0.86167 | 0.86377 | 0.89833 | 0.92 | 0.9182 | 1.0068
| Median |  | 0.841 | 1.0383 | 0.795 | 0.91925 | 0.9445 | 0.9 | 0.9715 | 0.86 | 0.86 | 0.95 | 0.93 | 0.87 | 0.895 | 0.87913 | 0.897 | 0.86 | 0.86509 | 0.9 | 0.92 | 0.9179 | 1.0067
| Std Dev. |  | 0.01002 | 0.022853 | 0.016021 | 0.0039423 | 0.0079162 | 0.021679 | 0.02315 | 0.0075277 | 0.01169 | 0.0075277 | 0.0075277 | 0.010328 | 0.019408 | 0.01551 | 0.012533 | 0.033714 | 0.012207 | 0.0040825 | 0.023664 | 0.011359 | 0.00033714
| Rel.Std.Dev. |  | 1.187% | 2.202% | 1.998% | 0.4296% | 0.8401% | 2.396% | 2.366% | 0.8736% | 1.362% | 0.7938% | 0.808% | 1.192% | 2.16% | 1.767% | 1.406% | 3.913% | 1.413% | 0.4545% | 2.572% | 1.237% | 0.03349%
| PDM3 |  | -6.18% | 15.38% | -10.89% | 2% | 4.751% | 0.6013% | 8.772% | -4.216% | -4.586% | 5.418% | 3.566% | -3.66% | -0.1398% | -2.405% | -0.918% | -4.216% | -3.982% | -0.1398% | 2.269% | 2.069% | 11.92%
| Z-Score (Absolute) |  | 1.06 | 2.72 | 1.89 | 0.38 | 0.86 | 0.13 | 1.56 | 0.72 | 0.78 | 0.98 | 0.65 | 0.62 | 0 | 0.4 | 0.14 | 0.72 | 0.67 | 0 | 0.42 | 0.39 | 2.12
| Analytical results for La in OREAS 235b (Certified Value 29.1 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 30.399 | 38.918 | 24.7 | 26.1433 | 35.47 | 26 | 36.9 | 29 | 27 | 26.3 | 29.7 | 29.2 | 27.9 | 28.079 | 19 | 29.74 | 30.4 | 34.9 | 24.2 | 23.8508 |  |  |  |  | 1
| 1 | 2 | 29.774 | 38.698 | 25.2 | 26.2633 | 35.46 | 22 | 36.4 | 29 | 28.4 | 27.1 | 29.4 | 29.7 | 30.2 | 28.1876 | 18 | 28.67 | 29.9 | 35.5 | 23.5 | 22.4797 |  |  |  |  | 28
| 1 | 3 | 30.325 | 37.851 | 24.9 | 26.9833 | 35.37 | 22 | 37.4 | 28 | 27.9 | 26.1 | 29.3 | 30.7 | 29.5 | 27.323 | 18 | 32.85 | 31 | 34.5 | 22.8 | 27.4047 |  |  |  |  | 16
| 1 | 4 | 30.403 | 38.786 | 24.9 | 26.2633 | 35.38 | 25 | 37.6 | 28 | 28.9 | 25.4 | 29.4 | 30.1 | 27.6 | 28.4051 | 20 | 30.09 | 31.3 | 36.3 | 22.4 | 27.1159 |  |  |  |  | 29.1008
| 1 | 5 | 28.654 | 38.808 | 25.8 | 26.2467 | 35.85 | 27 | 37.4 | 29 | 28.5 | 27.5 | 30.5 | 30.2 | 27.8 | 27.2259 | 19 | 28.53 | 33.3 | 35.2 | 22.5 | 26.7109 |  |  |  |  | 91
| 1 | 6 | 29.479 | 39.039 | 25.2 | 26.61 | 35.86 | 27 | 38.2 | 30 | 27.5 | 26.8 | 29.4 | 30.4 | 29.7 | 28.5254 | 18 | 31.69 | 31.3 | 34.8 | 24.4 | 23.2374
| Mean |  | 29.839 | 38.6833 | 25.1167 | 26.4183 | 35.565 | 24.8333 | 37.3167 | 28.8333 | 28.0333 | 26.5333 | 29.6167 | 30.05 | 28.7833 | 27.9577 | 18.6667 | 30.2617 | 31.2 | 35.2 | 23.3 | 25.1332
| Median |  | 30.0495 | 38.797 | 25.05 | 26.2633 | 35.465 | 25.5 | 37.4 | 29 | 28.15 | 26.55 | 29.4 | 30.15 | 28.7 | 28.1333 | 18.5 | 29.915 | 31.15 | 35.05 | 23.15 | 25.2809
| Std Dev. |  | 0.69368 | 0.42442 | 0.38687 | 0.31909 | 0.22828 | 2.3166 | 0.61455 | 0.75277 | 0.70333 | 0.75542 | 0.4535 | 0.53198 | 1.1409 | 0.55287 | 0.8165 | 1.7072 | 1.1662 | 0.63875 | 0.86718 | 2.1845
| Rel.Std.Dev. |  | 2.325% | 1.097% | 1.54% | 1.208% | 0.6419% | 9.329% | 1.647% | 2.611% | 2.509% | 2.847% | 1.531% | 1.77% | 3.964% | 1.978% | 4.374% | 5.642% | 3.738% | 1.815% | 3.722% | 8.692%
| PDM3 |  | 2.537% | 32.93% | -13.69% | -9.218% | 22.21% | -14.66% | 28.23% | -0.9192% | -3.668% | -8.823% | 1.773% | 3.262% | -1.091% | -3.928% | -35.86% | 3.989% | 7.213% | 20.96% | -19.93% | -13.63%
| Z-Score (Absolute) |  | 0.29 | 2.79 | 1.04 | 0.67 | 1.91 | 1.12 | 2.4 | 0.01 | 0.22 | 0.64 | 0.23 | 0.35 | 0.01 | 0.24 | 2.86 | 0.41 | 0.67 | 1.8 | 1.55 | 1.04
| Analytical results for Li in OREAS 235b (Certified Value 45.8 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 41.65 | 50.9882 | 50.6 | 48.175 | 45.4 | 49.1 | 52.8 | 46.4 | 44.3 | 43.3 | 44.7 | 45.3 | 46.4678 | 40.92 | 43 | 47.9 | 41.6 | 43.4329 | [6 empty] | 1
| 1 | 2 | 41.71 | 50.8426 | 50.1 | 48.13 | 44.2 | 46.1 | 51.4 | 46.1 | 46.5 | 43.5 | 43.4 | 45.7 | 44.3062 | 40.18 | 45 | 44.7 | 40.7 | 43.2363 | [6 empty] | 0
| 1 | 3 | 44.09 | 49.8115 | 47.7 | 48.22 | 45.3 | 47.6 | 52.9 | 46.9 | 46.7 | 43.4 | 43.9 | 46.2 | 48.2654 | 45.44 | 44 | 46.1 | 40.3 | 43.9035 | [6 empty] | 16
| 1 | 4 | 43.76 | 49.8159 | 51 | 47.815 | 44.4 | 48.1 | 52.9 | 46.2 | 48.1 | 43.6 | 44.1 | 47.1 | 43.3588 | 43.97 | 44 | 44.5 | 40.4 | 44.3546 | [6 empty] | 45.7841
| 1 | 5 | 42.84 | 49.6987 | 47.4 | 47.88 | 44.7 | 48.9 | 51.1 | 44.7 | 47.8 | 44.3 | 45.9 | 47.1 | 47.2132 | 42.59 | 41 | 43.2 | 41.3 | 41.3901 | [6 empty] | 92
| 1 | 6 | 43.41 | 48.6447 | 48 | 48.61 | 44.5 | 49.1 | 53.7 | 46.6 | 46.8 | 44.1 | 44.2 | 46.5 | 43.5722 | 42.31 | 42 | 45.6 | 41.9 | 43.8193
| Mean |  | 42.91 | 49.9669 | 49.1333 | 48.1383 | 44.75 | 48.15 | 52.4667 | 46.15 | 46.7 | 43.7 | 44.3667 | 46.3167 | 45.5306 | 42.5683 | 43.1667 | 45.3333 | 41.0333 | 43.3561
| Median |  | 43.125 | 49.8137 | 49.05 | 48.1525 | 44.6 | 48.5 | 52.85 | 46.3 | 46.75 | 43.55 | 44.15 | 46.35 | 45.387 | 42.45 | 43.5 | 45.15 | 41 | 43.6261
| Std Dev. |  | 1.039 | 0.85757 | 1.6071 | 0.28349 | 0.49295 | 1.1726 | 1.0013 | 0.76616 | 1.3401 | 0.40497 | 0.86178 | 0.73326 | 2.0611 | 1.9339 | 1.472 | 1.6058 | 0.66232 | 1.0389
| Rel.Std.Dev. |  | 2.421% | 1.716% | 3.271% | 0.5889% | 1.102% | 2.435% | 1.909% | 1.66% | 2.87% | 0.9267% | 1.942% | 1.583% | 4.527% | 4.543% | 3.41% | 3.542% | 1.614% | 2.396%
| PDM3 |  | -6.277% | 9.136% | 7.315% | 5.142% | -2.259% | 5.168% | 14.6% | 0.7992% | 2.001% | -4.552% | -3.096% | 1.163% | -0.5537% | -7.024% | -5.717% | -0.9845% | -10.38% | -5.303%
| Z-Score (Absolute) |  | 0.78 | 1.41 | 1.15 | 0.84 | 0.21 | 0.84 | 2.19 | 0.22 | 0.39 | 0.54 | 0.33 | 0.27 | 0.03 | 0.89 | 0.7 | 0.03 | 1.37 | 0.64
| Analytical results for Lu in OREAS 235b (Certified Value 0.13 ppm) | [25 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.21; 1.09: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.12913; 1.06: 0.1; 1.07: 0.113; 1.08: 0.12; 1.09: 0.1865 |  |  |  | 1
| 1 | 2 | 1.01: 0.13168; 1.06: 0.1; 1.07: 0.109; 1.08: 0.12; 1.09: 0.1712 |  |  |  | 29
| 1 | 3 | 1.01: 0.12211; 1.06: 0.1; 1.07: 0.123; 1.08: 0.13; 1.09: 0.1989 |  |  |  | 16
| 1 | 4 | 1.01: 0.12697; 1.06: 0.1; 1.07: 0.114; 1.08: 0.13; 1.09: 0.2063 |  |  |  | 0.12999
| 1 | 5 | 1.01: 0.12464; 1.06: 0.1; 1.07: 0.111; 1.08: 0.11; 1.09: 0.1971 |  |  |  | 93
| 1 | 6 | 1.01: 0.12629; 1.06: 0.1; 1.07: 0.119; 1.08: 0.11; 1.09: 0.1699
| Mean |  | 1.01: 0.1268; 1.06: 0.1; 1.07: 0.11483; 1.08: 0.12; 1.09: 0.18832
| Median |  | 1.01: 0.12663; 1.06: 0.1; 1.07: 0.1135; 1.08: 0.12; 1.09: 0.1918
| Std Dev. |  | 1.01: 0.0033504; 1.06: 1.5202e-17; 1.07: 0.0052313; 1.08: 0.0089443; 1.09: 0.015153
| Rel.Std.Dev. |  | 1.01: 2.642%; 1.06: 1.52e-14%; 1.07: 4.556%; 1.08: 7.454%; 1.09: 8.046%
| PDM3 |  | 1.01: -2.452%; 1.06: -23.07%; 1.07: -11.66%; 1.08: -7.686%; 1.09: 44.87%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.06: 1.98; 1.07: 0.51; 1.08: 0; 1.09: 6.77
| Analytical results for Mg in OREAS 235b (Certified Value 1.46 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.311 | 1.4165 | 1.45 | 1.4733 | 1.48 | 1.45 | 1.54 | 1.38 | 1.4 | 1.5 | 1.47 | 1.41 | 1.49 | 1.5971 | 1.532 | 1.35 | 1.395 | 1.5 | 1.55 | 1.6091 | 1.3769 |  |  |  | 1
| 1 | 2 | 1.293 | 1.4008 | 1.46 | 1.4733 | 1.48 | 1.4 | 1.53 | 1.38 | 1.43 | 1.48 | 1.47 | 1.39 | 1.54 | 1.5768 | 1.52 | 1.56 | 1.368 | 1.49 | 1.54 | 1.5604 | 1.5036 |  |  |  | 0
| 1 | 3 | 1.314 | 1.3478 | 1.41 | 1.49 | 1.44 | 1.42 | 1.54 | 1.38 | 1.42 | 1.48 | 1.48 | 1.4 | 1.5 | 1.5814 | 1.562 | 1.37 | 1.3849 | 1.49 | 1.48 | 1.5642 | 1.4817 |  |  |  | 16
| 1 | 4 | 1.306 | 1.4139 | 1.51 | 1.4567 | 1.45 | 1.43 | 1.56 | 1.37 | 1.44 | 1.46 | 1.49 | 1.41 | 1.5 | 1.5845 | 1.508 | 1.37 | 1.3642 | 1.49 | 1.46 | 1.5531 | 1.4911 |  |  |  | 1.4593
| 1 | 5 | 1.298 | 1.3886 | 1.41 | 1.4767 | 1.45 | 1.43 | 1.6 | 1.38 | 1.43 | 1.48 | 1.48 | 1.43 | 1.46 | 1.6015 | 1.556 | 1.3 | 1.4028 | 1.5 | 1.48 | 1.5923 | 1.489 |  |  |  | 94
| 1 | 6 | 1.284 | 1.3884 | 1.44 | 1.49 | 1.47 | 1.43 | 1.54 | 1.39 | 1.41 | 1.48 | 1.48 | 1.41 | 1.54 | 1.5785 | 1.526 | 1.32 | 1.3992 | 1.51 | 1.51 | 1.6197 | 1.5021
| Mean |  | 1.301 | 1.3926 | 1.4467 | 1.4767 | 1.4617 | 1.4267 | 1.5517 | 1.38 | 1.4217 | 1.48 | 1.4783 | 1.4083 | 1.505 | 1.5867 | 1.534 | 1.3783 | 1.3857 | 1.4967 | 1.5033 | 1.5831 | 1.4741
| Median |  | 1.302 | 1.3947 | 1.445 | 1.475 | 1.46 | 1.43 | 1.54 | 1.38 | 1.425 | 1.48 | 1.48 | 1.41 | 1.5 | 1.583 | 1.529 | 1.36 | 1.39 | 1.495 | 1.495 | 1.5782 | 1.4901
| Std Dev. |  | 0.011454 | 0.02503 | 0.037238 | 0.012472 | 0.017224 | 0.01633 | 0.025626 | 0.0063246 | 0.01472 | 0.012649 | 0.0075277 | 0.013292 | 0.030822 | 0.01024 | 0.021014 | 0.093256 | 0.016342 | 0.008165 | 0.036148 | 0.027831 | 0.048313
| Rel.Std.Dev. |  | 0.8804% | 1.797% | 2.574% | 0.8446% | 1.178% | 1.145% | 1.651% | 0.4583% | 1.035% | 0.8547% | 0.5092% | 0.9438% | 2.048% | 0.6454% | 1.37% | 6.766% | 1.179% | 0.5455% | 2.405% | 1.758% | 3.278%
| PDM3 |  | -10.85% | -4.568% | -0.8661% | 1.19% | 0.1618% | -2.237% | 6.329% | -5.434% | -2.579% | 1.418% | 1.304% | -3.493% | 3.131% | 8.727% | 5.119% | -5.549% | -5.045% | 2.56% | 3.017% | 8.485% | 1.012%
| Z-Score (Absolute) |  | 2.23 | 1.05 | 0.35 | 0.03 | 0.16 | 0.61 | 1 | 1.21 | 0.67 | 0.08 | 0.05 | 0.85 | 0.4 | 1.45 | 0.77 | 1.23 | 1.14 | 0.29 | 0.38 | 1.4 | 0
| Analytical results for Mn in OREAS 235b (Certified Value 0.031 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.03134 | 0.030277 | 0.0324 | 0.032248 | 0.0325 | 0.031 | 0.0303 | 0.0345 | 0.0292 | 0.0316 | 0.0303 | 0.0289 | 0.0299 | 0.032362 | 0.0279 | 0.031 | 0.028906 | 0.03459 | 0.0319 | 0.02902 | 0.031 |  |  |  | 1
| 1 | 2 | 0.03089 | 0.030681 | 0.0319 | 0.032078 | 0.0329 | 0.0299 | 0.0304 | 0.0338 | 0.0297 | 0.0312 | 0.0305 | 0.0282 | 0.0309 | 0.032141 | 0.0279 | 0.0317 | 0.028429 | 0.03429 | 0.0318 | 0.02839 | 0.0335 |  |  |  | 30
| 1 | 3 | 0.03171 | 0.029596 | 0.0315 | 0.032349 | 0.0321 | 0.0304 | 0.0305 | 0.0337 | 0.0298 | 0.0313 | 0.0307 | 0.0287 | 0.0304 | 0.032266 | 0.0287 | 0.0311 | 0.028707 | 0.03431 | 0.0302 | 0.028 | 0.0331 |  |  |  | 16
| 1 | 4 | 0.03178 | 0.030471 | 0.0324 | 0.032056 | 0.0322 | 0.031 | 0.0308 | 0.033 | 0.0301 | 0.031 | 0.0306 | 0.0286 | 0.0302 | 0.032079 | 0.0279 | 0.0311 | 0.028645 | 0.03405 | 0.0301 | 0.02773 | 0.0334 |  |  |  | 0.030969
| 1 | 5 | 0.03083 | 0.030183 | 0.0313 | 0.032007 | 0.0323 | 0.0311 | 0.0316 | 0.0339 | 0.03 | 0.0311 | 0.0302 | 0.0289 | 0.0295 | 0.032349 | 0.0287 | 0.0294 | 0.029022 | 0.0339 | 0.0306 | 0.02871 | 0.0331 |  |  |  | 95
| 1 | 6 | 0.03081 | 0.030321 | 0.0305 | 0.032381 | 0.033 | 0.0313 | 0.0307 | 0.0342 | 0.0296 | 0.031 | 0.0308 | 0.0286 | 0.0309 | 0.032231 | 0.0279 | 0.0299 | 0.028974 | 0.03412 | 0.0312 | 0.02902 | 0.0332
| Mean |  | 0.031227 | 0.030255 | 0.031667 | 0.032187 | 0.0325 | 0.030783 | 0.030717 | 0.03385 | 0.029733 | 0.0312 | 0.030517 | 0.02865 | 0.0303 | 0.032238 | 0.028167 | 0.0307 | 0.028781 | 0.03421 | 0.030967 | 0.028478 | 0.032883
| Median |  | 0.031115 | 0.030299 | 0.0317 | 0.032163 | 0.0324 | 0.031 | 0.0306 | 0.03385 | 0.02975 | 0.03115 | 0.03055 | 0.02865 | 0.0303 | 0.032248 | 0.0279 | 0.03105 | 0.028806 | 0.034205 | 0.0309 | 0.02855 | 0.03315
| Std Dev. |  | 0.00044653 | 0.00036668 | 0.00072847 | 0.00016055 | 0.00037417 | 0.00052694 | 0.00047081 | 0.00050892 | 0.00032042 | 0.00022804 | 0.00023166 | 0.00025884 | 0.00055498 | 0.0001126 | 0.00041312 | 0.00086487 | 0.00022752 | 0.00024108 | 0.00078655 | 0.00053611 | 0.00093684
| Rel.Std.Dev. |  | 1.43% | 1.212% | 2.3% | 0.4988% | 1.151% | 1.712% | 1.533% | 1.503% | 1.078% | 0.7309% | 0.7591% | 0.9035% | 1.832% | 0.3493% | 1.467% | 2.817% | 0.7905% | 0.7047% | 2.54% | 1.883% | 2.849%
| PDM3 |  | 0.8308% | -2.307% | 2.252% | 3.93% | 4.942% | -0.6007% | -0.816% | 9.302% | -3.991% | 0.7447% | -1.462% | -7.489% | -2.161% | 4.096% | -9.05% | -0.8698% | -7.068% | 10.46% | -0.008724% | -8.044% | 6.18%
| Z-Score (Absolute) |  | 0.28 | 0.34 | 0.57 | 0.9 | 1.1 | 0 | 0.04 | 1.97 | 0.67 | 0.27 | 0.17 | 1.37 | 0.31 | 0.93 | 1.7 | 0.05 | 1.29 | 2.2 | 0.12 | 1.48 | 1.35
| Analytical results for Mo in OREAS 235b (Certified Value 0.93 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.93 | 0.94 | < 2 | 1 | 0.65 | 1 | 1 | 0.95 | 1.14 | 0.87 | 0.91 | 0.94 | 0.96616 | < 1 | 1.038 | 1.05 | 0.55 | < 1 | 0.94 | 0.7249 | 2.5 |  |  |  | 1
| 1 | 2 | 0.9 | 0.95 | < 2 | 0.9 | 0.61 | 1 | 0.9 | 0.94 | 0.89 | 0.8 | 0.88 | 0.95 | 0.95877 | < 1 | 1.077 | 0.9 | 0.19 | < 1 | 0.98 | 0.5862 | 2.7168 |  |  |  | 31
| 1 | 3 | 0.92 | 0.88 | < 2 | 0.9 | 0.75 | 1.1 | 0.9 | 0.94 | 0.91 | 0.82 | 0.92 | 0.95 | 0.97955 | < 1 | 1.087 | 0.97 |  | < 1 | 0.98 | 0.7651 | 2.8253 |  |  |  | 16
| 1 | 4 | 0.91 | 0.93 | < 2 | 0.9 | 0.74 | 1 | 0.9 | 0.95 | 0.94 | 0.83 | 0.89 | 0.95 | 0.9518 | < 1 | 1.029 | 0.88 | 0.09 | < 1 | 1.07 | 0.9065 | 2.3924 |  |  |  | 0.93332
| 1 | 5 | 0.94 | 0.84 | < 2 | 0.9 | 0.72 | 1 | 0.9 | 0.91 | 0.92 | 0.85 | 0.9 | 0.85 | 0.92484 | < 1 | 1.043 | 0.83 |  | < 1 | 0.93 | 0.8481 | 2.501 |  |  |  | 96
| 1 | 6 | 0.9 | 0.93 | < 2 | 1 | 0.72 | 1.1 | 0.9 | 0.92 | 0.95 | 0.82 | 0.88 | 0.93 | 0.95177 | < 1 | 1 | 0.92 |  | < 1 | 0.91 | 0.7651 | 3.2594
| Mean |  | 0.91667 | 0.91167 |  | 0.93333 | 0.69833 | 1.0333 | 0.91667 | 0.935 | 0.95833 | 0.83167 | 0.89667 | 0.92833 | 0.95548 |  | 1.0457 | 0.925 | 0.27667 |  | 0.96833 | 0.76598 | 2.6991
| Median |  | 0.915 | 0.93 |  | 0.9 | 0.72 | 1 | 0.9 | 0.94 | 0.93 | 0.825 | 0.895 | 0.945 | 0.95528 |  | 1.0405 | 0.91 | 0.19 |  | 0.96 | 0.7651 | 2.6089
| Std Dev. |  | 0.01633 | 0.042622 |  | 0.05164 | 0.055648 | 0.05164 | 0.040825 | 0.016432 | 0.091524 | 0.024833 | 0.01633 | 0.0392 | 0.018277 |  | 0.031998 | 0.076616 | 0.24194 |  | 0.057067 | 0.11 | 0.31721
| Rel.Std.Dev. |  | 1.781% | 4.675% |  | 5.533% | 7.969% | 4.997% | 4.454% | 1.757% | 9.55% | 2.986% | 1.821% | 4.223% | 1.913% |  | 3.06% | 8.283% | 87.45% |  | 5.893% | 14.36% | 11.75%
| PDM3 |  | -1.784% | -2.32% |  | 0.001792% | -25.18% | 10.72% | -1.784% | 0.1804% | 2.68% | -10.89% | -3.927% | -0.5339% | 2.375% |  | 12.04% | -0.8911% | -70.36% |  | 3.752% | -17.93% | 189.2%
| Z-Score (Absolute) |  | 0.04 | 0.1 | 0.84 | NA | 2.37 | NA | NA | 0.15 | 0.4 | 0.95 | 0.26 | 0.08 | 0.37 | 4.48 | 1.33 | 0.04 | 6.86 | 4.48 | 0.51 | 1.65 | 18.93
Results from laboratories 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 235b (Certified Value 0.155 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.151 | 0.18453 | 0.161 | 0.13333 | 0.17 | 0.16 | 0.18 | 0.19 | 0.15 | 0.16 | 0.16 | 0.14 | 0.16 | 0.15065 | 0.141 | 0.15 | 0.15801 | 0.16 | 0.17 | 0.1442 | 0.0775 |  |  |  | 1
| 1 | 2 | 0.148 | 0.18554 | 0.15 | 0.13333 | 0.16 | 0.14 | 0.18 | 0.19 | 0.15 | 0.16 | 0.16 | 0.14 | 0.166 | 0.15628 | 0.141 | 0.15 | 0.15338 | 0.16 | 0.17 | 0.1408 | 0.068 |  |  |  | 0
| 1 | 3 | 0.151 | 0.18483 | 0.148 | 0.13 | 0.16 | 0.14 | 0.18 | 0.19 | 0.15 | 0.16 | 0.16 | 0.14 | 0.166 | 0.14896 | 0.141 | 0.15 | 0.15547 | 0.15 | 0.16 | 0.1393 | 0.067 |  |  |  | 16
| 1 | 4 | 0.149 | 0.18858 | 0.163 | 0.13 | 0.16 | 0.15 | 0.18 | 0.2 | 0.15 | 0.15 | 0.16 | 0.14 | 0.16 | 0.156 | 0.141 | 0.15 | 0.15509 | 0.16 | 0.16 | 0.1394 | 0.0692 |  |  |  | 0.15503
| 1 | 5 | 0.147 | 0.18198 | 0.148 | 0.13333 | 0.16 | 0.15 | 0.19 | 0.2 | 0.15 | 0.16 | 0.16 | 0.15 | 0.157 | 0.15131 | 0.148 | 0.14 | 0.15716 | 0.16 | 0.16 | 0.1424 | 0.0672 |  |  |  | 97
| 1 | 6 | 0.149 | 0.18739 | 0.151 | 0.13 | 0.17 | 0.15 | 0.18 | 0.21 | 0.15 | 0.16 | 0.16 | 0.14 | 0.166 | 0.14505 | 0.141 | 0.14 | 0.15921 | 0.16 | 0.16 | 0.144 | 0.0671
| Mean |  | 0.14917 | 0.18548 | 0.1535 | 0.13167 | 0.16333 | 0.14833 | 0.18167 | 0.19667 | 0.15 | 0.15833 | 0.16 | 0.14167 | 0.1625 | 0.15137 | 0.14217 | 0.14667 | 0.15638 | 0.15833 | 0.16333 | 0.14168 | 0.069333
| Median |  | 0.149 | 0.18519 | 0.1505 | 0.13167 | 0.16 | 0.15 | 0.18 | 0.195 | 0.15 | 0.16 | 0.16 | 0.14 | 0.163 | 0.15098 | 0.141 | 0.15 | 0.15631 | 0.16 | 0.16 | 0.1416 | 0.0676
| Std Dev. |  | 0.0016021 | 0.0023161 | 0.0067157 | 0.0018257 | 0.005164 | 0.0075277 | 0.0040825 | 0.008165 | 0 | 0.0040825 | 0 | 0.0040825 | 0.0039875 | 0.0042866 | 0.0028577 | 0.005164 | 0.0021316 | 0.0040825 | 0.005164 | 0.0021858 | 0.0040859
| Rel.Std.Dev. |  | 1.074% | 1.249% | 4.375% | 1.387% | 3.162% | 5.075% | 2.247% | 4.152% | 0% | 2.578% | 0% | 2.882% | 2.454% | 2.832% | 2.01% | 3.521% | 1.363% | 2.578% | 3.162% | 1.543% | 5.893%
| PDM3 |  | -3.783% | 19.64% | -0.988% | -15.07% | 5.355% | -4.321% | 17.18% | 26.86% | -3.246% | 2.13% | 3.205% | -8.621% | 4.817% | -2.359% | -8.298% | -5.396% | 0.8727% | 2.13% | 5.355% | -8.61% | -55.28%
| Z-Score (Absolute) |  | 0.32 | 2.4 | 0 | 1.64 | 0.74 | 0.39 | 2.11 | 3.23 | 0.26 | 0.36 | 0.49 | 0.89 | 0.67 | 0.16 | 0.85 | 0.51 | 0.22 | 0.36 | 0.74 | 0.89 | 6.31
| Analytical results for Nb in OREAS 235b (Certified Value 0.53 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.72 | 0.7 | < 5 | 0.9 | < 1 | 0.5 | 0.3 | 0.51 | 0.49 | 0.39 | 0.39 | 0.70661 | 0.74 | 0.4 | 0.4222 | [9 empty] | 1
| 1 | 2 | 0.77 | 0.7 | < 5 | 0.9 | < 1 | 0.5 | 0.3 | 0.49 | 0.56 | 0.37 | 0.45 | 0.72471 | 0.64 | 0.36 | 0.346 | [9 empty] | 32
| 1 | 3 | 0.74 | 0.8 | < 5 | 0.8 | < 1 | 0.6 | 0.4 | 0.44 | 0.49 | 0.41 | 0.45 | 0.76841 | 0.8 | 0.31 | 0.4201 | [9 empty] | 16
| 1 | 4 | 0.75 | 0.9 | < 5 | 0.8 | < 1 | 0.5 | 0.4 | 0.49 | 0.53 | 0.38 | 0.42 | 0.7437 | 0.65 | 0.39 | 0.4339 | [9 empty] | 0.52697
| 1 | 5 | 0.72 | 0.9 | < 5 | 0.8 | < 1 | 0.5 | 0.3 | 0.48 | 0.49 | 0.4 | 0.4 | 0.69851 | 0.6 | 0.35 | 0.4496 | [9 empty] | 98
| 1 | 6 | 0.74 | 0.9 | < 5 | 0.9 | < 1 | 0.5 | 0.5 | 0.49 | 0.51 | 0.36 | 0.46 | 0.7719 | 0.66 | 0.35 | 0.3789
| Mean |  | 0.74 | 0.81667 |  | 0.85 |  | 0.51667 | 0.36667 | 0.48333 | 0.51167 | 0.385 | 0.42833 | 0.73564 | 0.68167 | 0.36 | 0.40845
| Median |  | 0.74 | 0.85 |  | 0.85 |  | 0.5 | 0.35 | 0.49 | 0.5 | 0.385 | 0.435 | 0.7342 | 0.655 | 0.355 | 0.42115
| Std Dev. |  | 0.018974 | 0.098319 |  | 0.054772 |  | 0.040825 | 0.08165 | 0.023381 | 0.028577 | 0.018708 | 0.029269 | 0.030968 | 0.073869 | 0.032249 | 0.038572
| Rel.Std.Dev. |  | 2.564% | 12.04% |  | 6.444% |  | 7.902% | 22.27% | 4.837% | 5.585% | 4.859% | 6.833% | 4.21% | 10.84% | 8.958% | 9.444%
| PDM3 |  | 40.42% | 54.97% |  | 61.3% |  | -1.956% | -30.42% | -8.281% | -2.905% | -26.94% | -18.72% | 39.6% | 29.36% | -31.69% | -22.49%
| Z-Score (Absolute) |  | 1.41 | NA | 11.73 | NA | 0 | NA | NA | 0.1 | 0.07 | 0.67 | 0.42 | 1.38 | 1.07 | 0.82 | 0.54
| Analytical results for Nd in OREAS 235b (Indicative Value 24 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 1
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 1
| 1 | 1 | 1.01: 27.6743; 1.06: 21.5; 1.07: 25.54; 1.08: 19.9496 |  |  |  | 1
| 1 | 2 | 1.01: 28.1287; 1.06: 21.4; 1.07: 24.56; 1.08: 18.9161 |  |  |  | 9
| 1 | 3 | 1.01: 26.9768; 1.06: 21; 1.07: 28.3; 1.08: 23.0271 |  |  |  | 16
| 1 | 4 | 1.01: 27.197; 1.06: 21; 1.07: 25.62; 1.08: 22.9336 |  |  |  | 24.0205
| 1 | 5 | 1.01: 27.5861; 1.06: 22.4; 1.07: 24.46; 1.08: 22.713 |  |  |  | 15
| 1 | 6 | 1.01: 27.2648; 1.06: 21.3; 1.07: 27.16; 1.08: 19.884
| Mean |  | 1.01: 27.4713; 1.06: 21.4333; 1.07: 25.94; 1.08: 21.2372
| Median |  | 1.01: 27.4254; 1.06: 21.35; 1.07: 25.58; 1.08: 21.3313
| Std Dev. |  | 1.01: 0.41205; 1.06: 0.5164; 1.07: 1.511; 1.08: 1.8513
| Rel.Std.Dev. |  | 1.01: 1.5%; 1.06: 2.409%; 1.07: 5.825%; 1.08: 8.717%
| PDM3 |  | 1.01: 14.37%; 1.06: -10.77%; 1.07: 7.991%; 1.08: -11.59%
| Z-Score (Absolute) |  | 1.01: 1.09; 1.06: 0.65; 1.07: 0.65; 1.08: 0.7
| Analytical results for Ni in OREAS 235b (Certified Value 61 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 1
| 1 | 1 | 60.7 | 54.853 | 60.2 | 57.2233 | 62 | 67 | 67 | 67 | 60.8 | 63.8 | 65.4 | 60.1 | 62.6 | 67.2658 | 64 | 48.07 | 64.8 | 74.24 | 59.71 | 66.3 | 57.19 | 58.5599 | 62.776 |  | 1
| 1 | 2 | 60 | 54.8055 | 59.6 | 56.5667 | 63 | 65 | 66 | 66 | 62.1 | 63.1 | 63.2 | 60.3 | 63.9 | 64.9742 | 62 | 47.6 | 60.3 | 75.37 | 58.36 | 69.4 | 55.48 | 59.4951 | 64.0214 |  | 34
| 1 | 3 | 61.8 | 54.067 | 58 | 56.9767 | 61 | 65 | 66 | 65.4 | 62.5 | 62.9 | 63.6 | 60.5 | 62.2 | 64.0037 | 63 | 103 | 66.2 | 74.92 | 59.29 | 67.7 | 55.48 | 59.2137 | 62.9984 |  | 16
| 1 | 4 | 59.2 | 54.6795 | 61.2 | 56.53 | 62 | 66 | 66 | 65.4 | 63.1 | 60.9 | 64.9 | 60.4 | 63.2 | 64.6937 | 63 | 51.46 | 66.3 | 73.24 | 59.45 | 64.7 | 54.72 | 59.0275 | 63.7612 |  | 61.3819
| 1 | 5 | 62.1 | 53.6085 | 58.7 | 56.8967 | 61 | 65 | 68 | 66.3 | 63 | 64.5 | 67 | 61 | 61.9 | 64.2992 | 64 | 92.27 | 61.7 | 73.06 | 58.97 | 65.9 | 56.62 | 57.4942 | 64.289 |  | 99
| 1 | 6 | 60 | 54.3995 | 61.1 | 57.26 | 63 | 67 | 67 | 66.2 | 62.5 | 63.1 | 64.1 | 59.6 | 63.7 | 65.4385 | 63 | 46 | 66.5 | 75.68 | 59.01 | 64.5 | 57.475 | 55.9609 | 63.7378
| Mean |  | 60.6333 | 54.4022 | 59.8 | 56.9089 | 62 | 65.8333 | 66.6667 | 66.05 | 62.3333 | 63.05 | 64.7 | 60.3167 | 62.9167 | 65.1125 | 63.1667 | 64.7333 | 64.3 | 74.4183 | 59.1317 | 66.4167 | 56.1608 | 58.2919 | 63.5973
| Median |  | 60.35 | 54.5395 | 59.9 | 56.9367 | 62 | 65.5 | 66.5 | 66.1 | 62.5 | 63.1 | 64.5 | 60.35 | 62.9 | 64.834 | 63 | 49.765 | 65.5 | 74.58 | 59.15 | 66.1 | 56.05 | 58.7937 | 63.7495
| Std Dev. |  | 1.129 | 0.48685 | 1.2884 | 0.31227 | 0.89443 | 0.98319 | 0.8165 | 0.60581 | 0.83586 | 1.2095 | 1.3885 | 0.46224 | 0.81343 | 1.1686 | 0.75277 | 25.7719 | 2.6631 | 1.0968 | 0.46846 | 1.8681 | 1.0955 | 1.3402 | 0.5896
| Rel.Std.Dev. |  | 1.862% | 0.8949% | 2.155% | 0.5487% | 1.443% | 1.493% | 1.225% | 0.9172% | 1.341% | 1.918% | 2.146% | 0.7664% | 1.293% | 1.795% | 1.192% | 39.81% | 4.142% | 1.474% | 0.7922% | 2.813% | 1.951% | 2.299% | 0.9271%
| PDM3 |  | -1.219% | -11.37% | -2.577% | -7.287% | 1.007% | 7.252% | 8.61% | 7.605% | 1.55% | 2.718% | 5.406% | -1.735% | 2.5% | 6.078% | 2.908% | 5.46% | 4.754% | 21.24% | -3.666% | 8.202% | -8.506% | -5.034% | 3.609%
| Z-Score (Absolute) |  | 0.6 | 2.13 | 0.8 | 1.52 | 0.26 | 0.69 | 0.89 | 0.74 | 0.18 | 0 | 0.41 | 0.67 | 0.03 | 0.51 | 0.03 | 0.42 | 0.31 | 2.8 | 0.97 | 0.83 | 1.7 | 1.17 | 0.14
| Analytical results for P in OREAS 235b (Certified Value 0.068 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.0635 | 0.062 | 0.067185 | 0.068 | 0.07 | 0.0667 | 0.07 | 0.066 | 0.07 | 0.068 | 0.067 | 0.067 | 0.072276 | 0.0672 | 0.06 | 0.06769 | 0.07 | 0.0697 | 0.0753 | 0.075 |  |  |  |  | 1
| 1 | 2 | 0.0631 | 0.063 | 0.067341 | 0.068 | 0.0695 | 0.0668 | 0.07 | 0.066 | 0.069 | 0.067 | 0.067 | 0.068 | 0.069392 | 0.0668 | 0.07 | 0.066327 | 0.07 | 0.0693 | 0.0727 | 0.0702 |  |  |  |  | 0
| 1 | 3 | 0.0647 | 0.063 | 0.06756 | 0.067 | 0.0696 | 0.0675 | 0.069 | 0.066 | 0.07 | 0.07 | 0.068 | 0.067 | 0.069322 | 0.0681 | 0.06 | 0.06853 | 0.07 | 0.0679 | 0.0732 | 0.0698 |  |  |  |  | 16
| 1 | 4 | 0.0631 | 0.062 | 0.06744 | 0.067 | 0.0713 | 0.0667 | 0.068 | 0.067 | 0.068 | 0.069 | 0.067 | 0.066 | 0.069232 | 0.0672 | 0.06 | 0.066114 | 0.07 | 0.0662 | 0.0731 | 0.0706 |  |  |  |  | 0.067931
| 1 | 5 | 0.0645 | 0.06 | 0.067105 | 0.066 | 0.0702 | 0.0689 | 0.071 | 0.067 | 0.068 | 0.069 | 0.067 | 0.065 | 0.069246 | 0.0694 | 0.06 | 0.06691 | 0.07 | 0.0669 | 0.0739 | 0.0698 |  |  |  |  | 100
| 1 | 6 | 0.0643 | 0.059 | 0.067316 | 0.068 | 0.0708 | 0.0665 | 0.07 | 0.066 | 0.067 | 0.068 | 0.068 | 0.068 | 0.070298 | 0.0672 | 0.06 | 0.065952 | 0.07 | 0.0684 | 0.0767 | 0.0702
| Mean |  | 0.063867 | 0.0615 | 0.067324 | 0.067333 | 0.070233 | 0.067183 | 0.069667 | 0.066333 | 0.068667 | 0.0685 | 0.067333 | 0.066833 | 0.069961 | 0.06765 | 0.061667 | 0.06692 | 0.07 | 0.068067 | 0.07415 | 0.070933
| Median |  | 0.0639 | 0.062 | 0.067328 | 0.0675 | 0.0701 | 0.06675 | 0.07 | 0.066 | 0.0685 | 0.0685 | 0.067 | 0.067 | 0.069357 | 0.0672 | 0.06 | 0.066618 | 0.07 | 0.06815 | 0.07355 | 0.0702
| Std Dev. |  | 0.00072019 | 0.0016432 | 0.00016544 | 0.0008165 | 0.00070048 | 0.00090866 | 0.0010328 | 0.0005164 | 0.0012111 | 0.0010488 | 0.0005164 | 0.001169 | 0.0012039 | 0.00095864 | 0.0040825 | 0.0010123 | 0 | 0.0013545 | 0.0015489 | 0.0020146
| Rel.Std.Dev. |  | 1.128% | 2.672% | 0.2457% | 1.213% | 0.9974% | 1.353% | 1.482% | 0.7785% | 1.764% | 1.531% | 0.7669% | 1.749% | 1.721% | 1.417% | 6.62% | 1.513% | 0% | 1.99% | 2.089% | 2.84%
| PDM3 |  | -5.983% | -9.467% | -0.8927% | -0.8797% | 3.389% | -1.101% | 2.555% | -2.352% | 1.083% | 0.8377% | -0.8797% | -1.616% | 2.988% | -0.4135% | -9.222% | -1.487% | 3.046% | 0.1998% | 9.155% | 4.42%
| Z-Score (Absolute) |  | 2.09 | 3.46 | 0.09 | 0.09 | 1.59 | 0.17 | 1.26 | 0.67 | 0.68 | 0.59 | 0.09 | 0.38 | 1.43 | 0.09 | 3.36 | 0.33 | 1.45 | 0.34 | 3.85 | 1.99
| Analytical results for Pb in OREAS 235b (Certified Value 9.47 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  | 3
| 1 | 1 | 9.3 | 10.4 | 8.5433 | 10.3 | 9.6 | 9.3 | 9.5 | 9.8 | 9.2 | 9.7 | 9.3 | 9.07 | 10.0016 | 8 | 19.95 | 9.3 | 2.51 | 7.1 | 8.6 | 11.6 | 11 | 8.2908 |  |  | 1
| 1 | 2 | 9 | 10.4 | 9.1433 | 9.7 | 7.1 | 9.1 | 9.6 | 9.1 | 9 | 9.6 | 9.6 | 9.89 | 10.166 | 8 | 19.04 | 9.4 | 3.76 | 7.2 | 8 | 10.8 | 10 | 7.5302 |  |  | 35
| 1 | 3 | 9.4 | 10.2 | 9.0433 | 10.2 | 9.6 | 8.9 | 9.9 | 9.3 | 9.8 | 9.7 | 9.7 | 9.38 | 10.1053 | 8 | 21.76 | 9.5 |  | 7.1 | 8.6 | 10.8 | 11 | 7.3741 |  |  | 16
| 1 | 4 | 9.3 | 10.4 | 8.4667 | 9.7 | 8.2 | 9.4 | 9.7 | 9.4 | 9 | 10 | 9.5 | 9.53 | 9.2971 | 8 | 17.69 | 9.9 | 1.32 | 6.9 | 8.1 | 11.2 | 11 | 8.5323 |  |  | 9.4697
| 1 | 5 | 9.2 | 10 | 8.93 | 9.4 | 7.1 | 9 | 9.7 | 9.5 | 9.2 | 10.2 | 9.9 | 9.44 | 10.1527 | 9 | 22 | 9.3 | 0.76 | 6.9 | 7.9 | 10.6 | 12 | 6.5123 |  |  | 101
| 1 | 6 | 9.3 | 9.8 | 8.7267 | 9.7 | 7.4 | 8.9 | 9.1 | 9.2 | 9.1 | 9.7 | 9.6 | 11.76 | 9.6476 | 8 | 14.89 | 9.2 | 0.1 | 6.9 | 8.5 | 10.7 | 10 | 7.9532
| Mean |  | 9.25 | 10.2 | 8.8089 | 9.8333 | 8.1667 | 9.1 | 9.5833 | 9.3833 | 9.2167 | 9.8167 | 9.6 | 9.845 | 9.8951 | 8.1667 | 19.2217 | 9.4333 | 1.69 | 7.0167 | 8.2833 | 10.95 | 10.8333 | 7.6988
| Median |  | 9.3 | 10.3 | 8.8283 | 9.7 | 7.8 | 9.05 | 9.65 | 9.35 | 9.15 | 9.7 | 9.6 | 9.485 | 10.0534 | 8 | 19.495 | 9.35 | 1.32 | 7 | 8.3 | 10.8 | 11 | 7.7417
| Std Dev. |  | 0.13784 | 0.25298 | 0.27418 | 0.34448 | 1.181 | 0.20976 | 0.27142 | 0.24833 | 0.29944 | 0.23166 | 0.2 | 0.97461 | 0.35047 | 0.40825 | 2.6771 | 0.25033 | 1.4566 | 0.13292 | 0.31885 | 0.37815 | 0.75277 | 0.72822
| Rel.Std.Dev. |  | 1.49% | 2.48% | 3.113% | 3.503% | 14.46% | 2.305% | 2.832% | 2.646% | 3.249% | 2.36% | 2.083% | 9.9% | 3.542% | 4.999% | 13.93% | 2.654% | 86.19% | 1.894% | 3.849% | 3.453% | 6.949% | 9.459%
| PDM3 |  | -2.321% | 7.711% | -6.979% | 3.839% | -13.76% | -3.905% | 1.199% | -0.9125% | -2.673% | 3.663% | 1.375% | 3.963% | 4.491% | -13.76% | 103% | -0.3845% | -82.15% | -25.9% | -12.53% | 15.63% | 14.4% | -18.7%
| Z-Score (Absolute) |  | 0.23 | 1.16 | 0.88 | 0.62 | 1.81 | 0.45 | 0.26 | 0.04 | 0.28 | 0.6 | 0.28 | 0.64 | 0.71 | NA | 14.33 | 0.04 | 11.27 | 3.49 | 1.64 | 2.25 | NA | 2.5
Results from laboratories 1.19 and 1.31 were removed due to their 1 ppm reading resolution.
| Analytical results for Pd in OREAS 235b (Indicative Value 37.3 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.08; 1.07: 1.30 |  |  | ppb
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS |  |  | 1
|  |  | 1.01: 0.5g; 1.06: 01g; 1.07: 0.5g |  |  | 1
| 1 | 1 | 1.01: < 1; 1.06: < 10; 1.07: 39.5 |  |  | 1
| 1 | 2 | 1.01: < 1; 1.06: < 10; 1.07: 25.1 |  |  | 1
| 1 | 3 | 1.01: < 1; 1.06: < 10; 1.07: 38.6 |  |  | 16
| 1 | 4 | 1.01: < 1; 1.06: < 10; 1.07: 44.9 |  |  | 37.25
| 1 | 5 | 1.01: < 1; 1.06: < 10; 1.07: 36.8 |  |  | 16
| 1 | 6 | 1.01: < 1; 1.06: < 10; 1.07: 38.6
| Mean |  | 1.07: 37.25
| Median |  | 1.07: 38.6
| Std Dev. |  | 1.07: 6.5583
| Rel.Std.Dev. |  | 1.07: 17.61%
| PDM3 |  | 1.07: -2.22e-14%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.06: 0; 1.07: 4.83
| Analytical results for Pr in OREAS 235b (Indicative Value 6.48 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  | 2
| 1 | 1 | 1.01: 7.5637; 1.06: 5.9; 1.07: 6.915; 1.08: 5.3401 |  |  | 1
| 1 | 2 | 1.01: 7.3573; 1.06: 5.9; 1.07: 6.662; 1.08: 5.1071 |  |  | 11
| 1 | 3 | 1.01: 7.2579; 1.06: 5.8; 1.07: 7.698; 1.08: 6.2956 |  |  | 16
| 1 | 4 | 1.01: 7.3486; 1.06: 5.7; 1.07: 6.981; 1.08: 6.1595 |  |  | 6.4845
| 1 | 5 | 1.01: 7.2857; 1.06: 5.9; 1.07: 6.734; 1.08: 6.0875 |  |  | 17
| 1 | 6 | 1.01: 7.1915; 1.06: 5.8; 1.07: 7.365; 1.08: 5.2774
| Mean |  | 1.01: 7.3341; 1.06: 5.8333; 1.07: 7.0592; 1.08: 5.7112
| Median |  | 1.01: 7.3171; 1.06: 5.85; 1.07: 6.948; 1.08: 5.7138
| Std Dev. |  | 1.01: 0.12801; 1.06: 0.08165; 1.07: 0.398; 1.08: 0.52439
| Rel.Std.Dev. |  | 1.01: 1.745%; 1.06: 1.4%; 1.07: 5.638%; 1.08: 9.182%
| PDM3 |  | 1.01: 13.1%; 1.06: -10.04%; 1.07: 8.863%; 1.08: -11.92%
| Z-Score (Absolute) |  | 1.01: 0.89; 1.06: 0.61; 1.07: 0.61; 1.08: 0.74
| Analytical results for Pt in OREAS 235b (Indicative Value 2.3 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.08; 1.07: 1.30 |  |  | ppb
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 01g; 1.07: 0.5g |  |  | 2
| 1 | 1 | 1.01: 2; 1.06: < 5; 1.07: 2.4 |  |  | 1
| 1 | 2 | 1.01: 2; 1.06: < 5; 1.07: 1.6 |  |  | 1
| 1 | 3 | 1.01: 2; 1.06: < 5; 1.07: 2.6 |  |  | 16
| 1 | 4 | 1.01: 2; 1.06: < 5; 1.07: 2.7 |  |  | 2.3
| 1 | 5 | 1.01: 2; 1.06: < 5; 1.07: 3.9 |  |  | 18
| 1 | 6 | 1.01: 2; 1.06: < 5; 1.07: 2.4
| Mean |  | 1.01: 2; 1.07: 2.6
| Median |  | 1.01: 2; 1.07: 2.5
| Std Dev. |  | 1.01: 0; 1.07: 0.74565
| Rel.Std.Dev. |  | 1.01: 0%; 1.07: 28.68%
| PDM3 |  | 1.01: -13.04%; 1.07: 13.04%
| Z-Score (Absolute) |  | 1.01: 3.37; 1.06: 0; 1.07: 0.67
| Analytical results for Rb in OREAS 235b (Certified Value 88 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 87.79 | 70.6245 | 68.3 | 88.62 | 84.8 | 101 | 83.6 | 87.4 | 88.1 | 91.4 | 86.4 | 86.1829 | 78.47 | 91.9 | 92.7 | 87.4995 | 89.6455 | [7 empty] | 1
| 1 | 2 | 86.35 | 70.5118 | 68.3 | 90.07 | 81.9 | 99.3 | 83.2 | 90.1 | 88.5 | 88.8 | 86.9 | 87.74 | 74.31 | 88.6 | 95.8 | 80.1342 | 87.7991 | [7 empty] | 37
| 1 | 3 | 89.91 | 69.6249 | 66.8 | 89.06 | 81.6 | 100 | 82.3 | 89.1 | 85.7 | 87.8 | 86.3 | 86.4653 | 82.18 | 92 | 93.3 | 97.626 | 89.1685 | [7 empty] | 16
| 1 | 4 | 89.7 | 69.9209 | 70 | 89.22 | 82.1 | 100 | 82 | 90.7 | 84.1 | 89.3 | 88.2 | 87.3962 | 78.9 | 92.6 | 88.6 | 96.065 | 90.95 | [7 empty] | 88.1365
| 1 | 5 | 88.1 | 70.0802 | 67.8 | 88.88 | 83.9 | 101 | 82.7 | 89.9 | 89.4 | 92.7 | 88.8 | 87.0234 | 76.21 | 87.2 | 91.1 | 96.4555 | 87.7791 | [7 empty] | 102
| 1 | 6 | 88.6 | 70.1827 | 66.3 | 89.99 | 85.7 | 104 | 82.9 | 89.9 | 86.3 | 89 | 85.9 | 86.6339 | 73.75 | 87.4 | 89.4 | 85.2621 | 90.5299
| Mean |  | 88.4083 | 70.1575 | 67.9167 | 89.3067 | 83.3333 | 100.8833 | 82.7833 | 89.5167 | 87.0167 | 89.8333 | 87.0833 | 86.907 | 77.3033 | 89.95 | 91.8167 | 90.5071 | 89.312
| Median |  | 88.35 | 70.1315 | 68.05 | 89.14 | 83 | 100.5 | 82.8 | 89.9 | 87.2 | 89.15 | 86.65 | 86.8286 | 77.34 | 90.25 | 91.9 | 91.7823 | 89.407
| Std Dev. |  | 1.3171 | 0.37144 | 1.3045 | 0.59524 | 1.7119 | 1.6618 | 0.58452 | 1.1566 | 1.9904 | 1.8359 | 1.1583 | 0.59004 | 3.1772 | 2.4866 | 2.6664 | 7.2264 | 1.3369
| Rel.Std.Dev. |  | 1.49% | 0.5294% | 1.921% | 0.6665% | 2.054% | 1.647% | 0.7061% | 1.292% | 2.287% | 2.044% | 1.33% | 0.6789% | 4.11% | 2.764% | 2.904% | 7.984% | 1.497%
| PDM3 |  | 0.3084% | -20.4% | -22.94% | 1.328% | -5.45% | 14.46% | -6.074% | 1.566% | -1.271% | 1.925% | -1.195% | -1.395% | -12.29% | 2.058% | 4.176% | 2.69% | 1.334%
| Z-Score (Absolute) |  | 0 | 7.98 | 8.96 | 0.39 | 2.22 | 5.46 | 2.46 | 0.48 | 0.61 | 0.62 | 0.58 | 0.66 | 4.86 | 0.67 | 1.49 | 0.92 | 0.4
| Analytical results for Re in OREAS 235b (Certified Value < 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.06; 1.07: 1.08; 1.08: 1.09; 1.09: 1.10; 1.10: 1.11; 1.11: 1.12; 1.12: 1.13; 1.13: 1.14; 1.14: 1.15; 1.15: 1.27; 1.16: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.09: AR*OES/MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*MS; 1.15: AR*OES/MS; 1.16: AR*MS |  |  | 3
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 01g; 1.08: 0.2g; 1.09: 01g; 1.10: 0.25g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g; 1.14: 0.5g; 1.15: 0.5g; 1.16: 0.5g |  |  | 3
| 1 | 1 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: 0.001; 1.12: 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0003 |  |  | 1
| 1 | 2 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0007 |  |  | 38
| 1 | 3 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.001 |  |  | 16
| 1 | 4 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0008 |  |  | < 0.001
| 1 | 5 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0007 |  |  | 103
| 1 | 6 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0015
| Mean |  | 1.11: 0.001; 1.12: 0.001; 1.16: 0.00083333
| Median |  | 1.11: 0.001; 1.12: 0.001; 1.16: 0.00075
| Std Dev. |  | 1.11: 0; 1.16: 0.00039833
| Rel.Std.Dev. |  | 1.11: 0%; 1.16: 47.8%
PDM3
| Z-Score (Absolute) |  | 1.01: 0.67; 1.06: 0.67; 1.07: 131.49; 1.08: 2.02; 1.09: 131.49; 1.10: 10.11; 1.11: 0.22; 1.12: 0.22; 1.13: 0.67; 1.14: 0.67; 1.15: 0.67; 1.16: 1.12
| Analytical results for Rh in OREAS 235b (Indicative Value 0.001 ppm) | [25 empty] | Indicative
Rh | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 3
|  |  | 1.01: 0.5g |  |  | 3
| 1 | 1 | 1.01: 0.0002 |  |  | 1
| 1 | 2 | 1.01: 0.0005 |  |  | 13
| 1 | 3 | 1.01: 0.0005 |  |  | 16
| 1 | 4 | 1.01: 0.0014 |  |  | 0.00083333
| 1 | 5 | 1.01: 0.0013 |  |  | 19
| 1 | 6 | 1.01: 0.0011
| Mean |  | 1.01: 0.00083333
| Median |  | 1.01: 0.0008
| Std Dev. |  | 1.01: 0.00049666
| Rel.Std.Dev. |  | 1.01: 59.6%
| PDM3 |  | 1.01: 4.441e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ru in OREAS 235b (Indicative Value < 0.0065 ppm) | [25 empty] | Indicative
Ru | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 3
|  |  | 1.01: 0.5g |  |  | 3
| 1 | 1 | 1.01: < 0.0009 |  |  | 1
| 1 | 2 | 1.01: < 0.0019 |  |  | 14
| 1 | 3 | 1.01: < 0.0047 |  |  | 16
| 1 | 4 | 1.01: < 0.0065 |  |  | < 0.0065
| 1 | 5 | 1.01: < 0.0065 |  |  | 20
| 1 | 6 | 1.01: 0.0009
| Mean |  | 1.01: 0.0009
| Median |  | 1.01: 0.0009
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for S in OREAS 235b (Certified Value 0.143 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.13 | 0.15251 | 0.133 | 0.12 | 0.15 | 0.13 | 0.14 | 0.15 | 0.15 | 0.16 | 0.15 | 0.13 | 0.15 | 0.122 | 0.13 | 0.006504 | 0.15 | 0.15 | 0.1673 | 0.1404 |  |  |  |  | 1
| 1 | 2 | 0.13 | 0.15263 | 0.135 | 0.12 | 0.15 | 0.13 | 0.14 | 0.15 | 0.16 | 0.15 | 0.15 | 0.13 | 0.15 | 0.121 | 0.14 | 0.006556 | 0.15 | 0.15 | 0.1635 | 0.6434 |  |  |  |  | 23
| 1 | 3 | 0.14 | 0.14947 | 0.137 | 0.12267 | 0.15 | 0.13 | 0.14 | 0.16 | 0.16 | 0.15 | 0.15 | 0.13 | 0.15 | 0.12 | 0.13 | 0.006405 | 0.14 | 0.15 | 0.1618 | 0.1704 |  |  |  |  | 16
| 1 | 4 | 0.13 | 0.15597 | 0.141 | 0.12 | 0.15 | 0.13 | 0.15 | 0.15 | 0.16 | 0.15 | 0.15 | 0.13 | 0.15 | 0.12 | 0.14 | 0.006216 | 0.15 | 0.14 | 0.1599 | 0.2624 |  |  |  |  | 0.14348
| 1 | 5 | 0.14 | 0.14978 | 0.13 | 0.12267 | 0.15 | 0.13 | 0.15 | 0.15 | 0.16 | 0.15 | 0.15 | 0.14 | 0.14 | 0.122 | 0.13 | 0.006395 | 0.15 | 0.15 | 0.1665 | 0.1483 |  |  |  |  | 104
| 1 | 6 | 0.14 | 0.14884 | 0.13 | 0.12 | 0.15 | 0.13 | 0.14 | 0.16 | 0.16 | 0.15 | 0.15 | 0.14 | 0.15 | 0.12 | 0.13 | 0.006252 | 0.15 | 0.15 | 0.1697 | 0.1428
| Mean |  | 0.135 | 0.15153 | 0.13433 | 0.12089 | 0.15 | 0.13 | 0.14333 | 0.15333 | 0.15833 | 0.15167 | 0.15 | 0.13333 | 0.14833 | 0.12083 | 0.13333 | 0.006388 | 0.14833 | 0.14833 | 0.16478 | 0.25128
| Median |  | 0.135 | 0.15114 | 0.134 | 0.12 | 0.15 | 0.13 | 0.14 | 0.15 | 0.16 | 0.15 | 0.15 | 0.13 | 0.15 | 0.1205 | 0.13 | 0.0064 | 0.15 | 0.15 | 0.165 | 0.15935
| Std Dev. |  | 0.0054772 | 0.0026992 | 0.004274 | 0.0013771 | 0 | 0 | 0.005164 | 0.005164 | 0.0040825 | 0.0040825 | 0 | 0.005164 | 0.0040825 | 0.00098319 | 0.005164 | 0.00013424 | 0.0040825 | 0.0040825 | 0.0036837 | 0.19753
| Rel.Std.Dev. |  | 4.057% | 1.781% | 3.182% | 1.139% | 0% | 0% | 3.603% | 3.368% | 2.578% | 2.692% | 0% | 3.873% | 2.752% | 0.8137% | 3.873% | 2.101% | 2.752% | 2.752% | 2.235% | 78.61%
| PDM3 |  | -5.912% | 5.61% | -6.377% | -15.75% | 4.542% | -9.397% | -0.1044% | 6.865% | 10.35% | 5.703% | 4.542% | -7.074% | 3.38% | -15.79% | -7.074% | -95.55% | 3.38% | 3.38% | 14.85% | 75.13%
| Z-Score (Absolute) |  | 0.77 | 0.18 | 0.81 | 1.59 | 0.1 | 1.06 | 0.29 | 0.29 | 0.58 | 0.19 | 0.1 | 0.87 | 0 | 1.59 | 0.87 | 8.2 | 0 | 0 | 0.95 | 5.95
| Analytical results for Sb in OREAS 235b (Certified Value 238 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 0
| 1 | 1 | 281.72 | 205 | 257.9233 | 215.23 | 158.07 | 213 | 250.29 | 249 | 283 | 266 | 255 | 291.83 | 287.1267 | 207 | 275 | 207 | 307.16 | 295.1 | 224.75 | 184.1 | 161.7595 |  |  |  | 1
| 1 | 2 | 274.36 | 197 | 259.79 | 219.23 | 142.21 | 209 | 236.95 | 257 | 280 | 257 | 255 | 301.14 | 289.3692 | 207 | 215.9 | 208 | 298 | 297.4 | 218.87 | 178.8 | 161.8178 |  |  |  | 39
| 1 | 3 | 280.88 | 192 | 261.595 | 203.16 | 139.54 | 213 | 257.2 | 250 | 285 | 268 | 263 | 287.59 | 285.3649 | 200 | 285.2 | 216 | 311.17 | 294.9 | 240.14 | 179.5 | 161.8916 |  |  |  | 16
| 1 | 4 | 286.49 | 178 | 260.605 | 196.13 | 144.92 | 227 | 239.9 | 258 | 281 | 272 | 252 | 323.72 | 290.6662 | 203 | 255.7 | 210 | 304.19 | 297.5 | 212.73 | 174.8 | 161.9285 |  |  |  | 238.2006
| 1 | 5 | 279.51 | 186 | 265.15 | 199.6 | 154.41 | 199 | 240.12 | 253 | 271 | 267 | 263 | 287.74 | 291.7973 | 205 | 224.7 | 198 | 301.28 | 296.7 | 221.99 | 182.2 | 161.9285 |  |  |  | 105
| 1 | 6 | 277.63 | 184 | 263.5033 | 202.45 | 155.26 | 209 | 243.13 | 252 | 282 | 264 | 265 | 292.64 | 285.393 | 200 | 223.9 | 208 | 295.81 | 295 | 219.39 | 180.9 | 161.7595
| Mean |  | 280.0983 | 190.3333 | 261.4278 | 205.9667 | 149.0683 | 211.6667 | 244.5983 | 253.1667 | 280.3333 | 265.6667 | 258.8333 | 297.4433 | 288.2862 | 203.6667 | 246.7333 | 207.8333 | 302.935 | 296.1 | 222.9783 | 180.05 | 161.8476
| Median |  | 280.195 | 189 | 261.1 | 202.805 | 149.665 | 211 | 241.625 | 252.5 | 281.5 | 266.5 | 259 | 292.235 | 288.248 | 204 | 240.2 | 208 | 302.735 | 295.9 | 220.69 | 180.2 | 161.8547
| Std Dev. |  | 4.0856 | 9.7297 | 2.6012 | 9.1571 | 7.7838 | 9.0921 | 7.6679 | 3.656 | 4.8854 | 5.0067 | 5.4559 | 13.7837 | 2.7355 | 3.2042 | 29.3824 | 5.8109 | 5.7463 | 1.2377 | 9.3104 | 3.1992 | 0.079293
| Rel.Std.Dev. |  | 1.459% | 5.112% | 0.995% | 4.446% | 5.222% | 4.295% | 3.135% | 1.444% | 1.743% | 1.885% | 2.108% | 4.634% | 0.9489% | 1.573% | 11.91% | 2.796% | 1.897% | 0.418% | 4.175% | 1.777% | 0.04899%
| PDM3 |  | 17.59% | -20.1% | 9.751% | -13.53% | -37.42% | -11.14% | 2.686% | 6.283% | 17.69% | 11.53% | 8.662% | 24.87% | 21.03% | -14.5% | 3.582% | -12.75% | 27.18% | 24.31% | -6.391% | -24.41% | -32.05%
| Z-Score (Absolute) |  | 0.58 | 0.98 | 0.25 | 0.71 | 1.69 | 0.61 | 0.04 | 0.11 | 0.58 | 0.33 | 0.21 | 0.88 | 0.72 | 0.75 | 0 | 0.67 | 0.97 | 0.86 | 0.41 | 1.16 | 1.47
| Analytical results for Sc in OREAS 235b (Certified Value 7.82 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 7.9 | 6.7617 | 7.3 | 7.3767 | 8 | 9.1 | 8 | 7.5 | 8.4 | 8.7 | 8.5 | 8.3 | 7.2 | 8.524 | 7 | 7.083 | 8.2 | 7.8 | 7.8 | 7.39 | 10.4911 | 7.7662 |  |  | 1
| 1 | 2 | 7.98 | 6.6778 | 7.2 | 7.3033 | 8 | 8.5 | 8 | 7.5 | 8.7 | 8.8 | 8.2 | 8.4 | 7.4 | 8.4036 | 7 | 6.942 | 7.7 | 7.5 | 7.3 | 7.21 | 11.1082 | 7.6923 |  |  | 40
| 1 | 3 | 7.93 | 6.5662 | 7.3 | 7.3833 | 7 | 8.6 | 8 | 7.4 | 8.7 | 8.5 | 8.2 | 8.4 | 7.3 | 8.5729 | 7 | 7.242 | 8.8 | 7.7 | 7.5 | 7.2 | 11.1093 | 7.9149 |  |  | 16
| 1 | 4 | 7.91 | 6.4248 | 7.6 | 7.3133 | 8 | 8.8 | 9 | 7.4 | 8.9 | 8.4 | 8.4 | 8.4 | 7.2 | 8.9498 | 7 | 7.239 | 8.4 | 7.7 | 7.3 | 7.15 | 11.1104 | 7.9105 |  |  | 7.8245
| 1 | 5 | 7.91 | 6.5964 | 7 | 7.2833 | 7 | 8.7 | 9 | 7.5 | 8.8 | 8.7 | 8.4 | 8.6 | 7 | 8.1681 | 7 | 6.992 | 7.9 | 7.7 | 7.1 | 7.31 | 10.9883 | 8.0168 |  |  | 106
| 1 | 6 | 7.86 | 6.6871 | 7.2 | 7.4167 | 8 | 9.1 | 9 | 7.5 | 8.8 | 8.7 | 8.3 | 8.2 | 7.4 | 8.5777 | 7 | 7.5 | 8.1 | 7.7 | 7.5 | 7.44 | 11.2359 | 7.7708
| Mean |  | 7.915 | 6.619 | 7.2667 | 7.3461 | 7.6667 | 8.8 | 8.5 | 7.4667 | 8.7167 | 8.6333 | 8.3333 | 8.3833 | 7.25 | 8.5327 | 7 | 7.1663 | 8.1833 | 7.6833 | 7.4167 | 7.2833 | 11.0072 | 7.8452
| Median |  | 7.91 | 6.6371 | 7.25 | 7.345 | 8 | 8.75 | 8.5 | 7.5 | 8.75 | 8.7 | 8.35 | 8.4 | 7.25 | 8.5485 | 7 | 7.161 | 8.15 | 7.7 | 7.4 | 7.26 | 11.1088 | 7.8406
| Std Dev. |  | 0.03937 | 0.11787 | 0.19664 | 0.053184 | 0.5164 | 0.25298 | 0.54772 | 0.05164 | 0.17224 | 0.15055 | 0.12111 | 0.13292 | 0.15166 | 0.25586 | 0 | 0.20494 | 0.38687 | 0.098319 | 0.24014 | 0.11553 | 0.26469 | 0.12144
| Rel.Std.Dev. |  | 0.4974% | 1.781% | 2.706% | 0.724% | 6.736% | 2.875% | 6.444% | 0.6916% | 1.976% | 1.744% | 1.453% | 1.585% | 2.092% | 2.999% | 0% | 2.86% | 4.728% | 1.28% | 3.238% | 1.586% | 2.405% | 1.548%
| PDM3 |  | 1.157% | -15.41% | -7.129% | -6.114% | -2.017% | 12.47% | 8.633% | -4.573% | 11.4% | 10.34% | 6.503% | 7.142% | -7.342% | 9.051% | -10.54% | -8.412% | 4.586% | -1.804% | -5.212% | -6.916% | 40.68% | 0.265%
| Z-Score (Absolute) |  | 0.08 | 1.47 | 0.69 | 0.6 | NA | 1.15 | NA | 0.45 | 1.05 | 0.95 | 0.59 | 0.65 | 0.71 | 0.82 | NA | 0.81 | 0.41 | 0.19 | 0.51 | 0.67 | 3.79 | 0
Results from laboratories 1.08, 1.10 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 235b (Indicative Value 0.23 ppm) | [25 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.15 | 0.3 | < 5 | < 1 | 0.7 | < 1 | 2.2 | 0.3 | 0.2 | 0.2 | 0.2 | 0.1 | < 1 | 0.7 | < 1 | < 0.2 | 0.1876 | 3.2008 | 0.30495 | [5 empty] | 1
| 1 | 2 | 0.14 | < 0.1 | < 5 | < 1 | 0.7 | < 1 | 2.2 | < 0.2 | 0.2 | < 0.2 | 0.3 | 0.1 | < 1 | 0.8 | < 1 | < 0.2 | 0.1504 | 2.4446 | 0.32291 | [5 empty] | 41
| 1 | 3 | 0.15 | 0.2 | < 5 | < 1 | 0.7 | < 1 | 2.4 | 0.2 | 0.2 | < 0.2 | 0.2 | < 0.1 | < 1 | 0.9 | < 1 | < 0.2 | 0.1738 | 2.5874 | 0.32447 | [5 empty] | 16
| 1 | 4 | 0.15 | < 0.1 | < 5 | < 1 | 0.7 | < 1 | 2.1 | 0.2 | 0.2 | 0.3 | 0.2 | < 0.1 | < 1 | 0.1 | < 1 | < 0.2 | 0.2138 | 2.6315 | 0.33005 | [5 empty] | 0.23465
| 1 | 5 | 0.15 | < 0.1 | < 5 | < 1 | 0.6 | < 1 | 2.1 | 0.3 | 0.4 | 0.2 | 0.3 | 0.2 | < 1 | 0.11 | < 1 | < 0.2 | 0.2359 | 3.0595 | 0.31491 | [5 empty] | 14
| 1 | 6 | 0.14 | 0.4 | < 5 | < 1 | 0.6 | < 1 | 2.3 | 0.2 | 0.3 | 0.3 | 0.4 | 0.1 | < 1 | 0.12 | < 1 | < 0.2 | 0.178 | 2.9611 | 0.32735
| Mean |  | 0.14667 | 0.3 |  |  | 0.66667 |  | 2.2167 | 0.24 | 0.25 | 0.25 | 0.26667 | 0.125 |  | 0.455 |  |  | 0.18992 | 2.8142 | 0.32077
| Median |  | 0.15 | 0.3 |  |  | 0.7 |  | 2.2 | 0.2 | 0.2 | 0.25 | 0.25 | 0.1 |  | 0.41 |  |  | 0.1828 | 2.7963 | 0.32369
| Std Dev. |  | 0.005164 | 0.1 |  |  | 0.05164 |  | 0.1169 | 0.054772 | 0.083666 | 0.057735 | 0.08165 | 0.05 |  | 0.38324 |  |  | 0.030508 | 0.30088 | 0.0092983
| Rel.Std.Dev. |  | 3.521% | 33.33% |  |  | 7.746% |  | 5.274% | 22.82% | 33.47% | 23.09% | 30.62% | 40% |  | 84.23% |  |  | 16.06% | 10.69% | 2.899%
| PDM3 |  | -37.5% | 27.85% |  |  | 184.1% |  | 844.7% | 2.279% | 6.541% | 6.541% | 13.64% | -46.73% |  | 93.9% |  |  | -19.06% | 1099% | 36.7%
| Z-Score (Absolute) |  | 0.66 | 0.55 | 8.2 | 0.67 | 1.3 | 0.67 | 7.13 | 0.39 | 0.27 | 0.45 | 0.2 | 0.83 | 0.67 | 0.51 | 0.67 | 0.83 | 0.49 | 9.38 | 0
| Analytical results for Si in OREAS 235b (Indicative Value 0.477 wt.%) | [25 empty] | Indicative
Si | Round | Replicate | 1.01: Lab |  |  |  | 1
| No | No | 1.01: 1.30 |  |  |  | wt.%
|  |  | 1.01: AR*OES |  |  |  | 3
|  |  | 1.01: 0.5g |  |  |  | 3
| 1 | 1 | 1.01: 0.492 |  |  |  | 1
| 1 | 2 | 1.01: 0.4924 |  |  |  | 9
| 1 | 3 | 1.01: 0.4603 |  |  |  | 16
| 1 | 4 | 1.01: 0.4732 |  |  |  | 0.47685
| 1 | 5 | 1.01: 0.4647 |  |  |  | 15
| 1 | 6 | 1.01: 0.4785
| Mean |  | 1.01: 0.47685
| Median |  | 1.01: 0.47585
| Std Dev. |  | 1.01: 0.013483
| Rel.Std.Dev. |  | 1.01: 2.827%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sm in OREAS 235b (Indicative Value 4.42 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 5.2742; 1.06: 3.1; 1.07: 4.693; 1.08: 3.8542 |  |  |  | 1
| 1 | 2 | 1.01: 5.1345; 1.06: 3.7; 1.07: 4.506; 1.08: 3.7589 |  |  |  | 10
| 1 | 3 | 1.01: 5.1289; 1.06: 3.7; 1.07: 5.135; 1.08: 4.4664 |  |  |  | 16
| 1 | 4 | 1.01: 5.1264; 1.06: 4; 1.07: 4.698; 1.08: 4.4471 |  |  |  | 4.4205
| 1 | 5 | 1.01: 5.1052; 1.06: 3.4; 1.07: 4.499; 1.08: 4.4454 |  |  |  | 16
| 1 | 6 | 1.01: 5.1468; 1.06: 4.1; 1.07: 4.954; 1.08: 3.7186
| Mean |  | 1.01: 5.1527; 1.06: 3.6667; 1.07: 4.7475; 1.08: 4.1151
| Median |  | 1.01: 5.1317; 1.06: 3.7; 1.07: 4.6955; 1.08: 4.1498
| Std Dev. |  | 1.01: 0.061069; 1.06: 0.37238; 1.07: 0.25231; 1.08: 0.3728
| Rel.Std.Dev. |  | 1.01: 1.185%; 1.06: 10.16%; 1.07: 5.315%; 1.08: 9.059%
| PDM3 |  | 1.01: 16.56%; 1.06: -17.05%; 1.07: 7.398%; 1.08: -6.908%
| Z-Score (Absolute) |  | 1.01: 0.94; 1.06: 0.99; 1.07: 0.41; 1.08: 0.41
| Analytical results for Sn in OREAS 235b (Certified Value 1.76 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.76 | 2.5324 | 1.71 | < 10 | 1.7 | 1.9 | 2.1 | 1.7 | 1.7 | 1.7 | 1.7 | 1.8 | 1.8716 | < 1 | 1.8 | < 5 | 1.7 | 1.4205 | 2.9175 | [5 empty] | 1
| 1 | 2 | 1.73 | 2.614 | 1.64 | < 10 | 1.7 | 1.8 | 1.9 | 1.8 | 1.8 | 1.7 | 1.7 | 1.8 | 1.8885 | < 1 | 1.8 | < 5 | 1.8 | 1.3129 | 2.6024 | [5 empty] | 26
| 1 | 3 | 1.84 | 2.551 | 1.67 | < 10 | 1.6 | 1.8 | 1.9 | 1.7 | 1.7 | 1.7 | 1.8 | 1.8 | 1.8549 | < 1 | 1.9 | < 5 | 1.8 | 1.6308 | 2.6172 | [5 empty] | 16
| 1 | 4 | 1.79 | 2.574 | 1.6 | < 10 | 1.7 | 1.9 | 2 | 1.7 | 1.8 | 1.7 | 1.8 | 1.8 | 1.8751 | < 1 | 1.8 | < 5 | 1.7 | 1.6357 | 2.6741 | [5 empty] | 1.7558
| 1 | 5 | 1.79 | 2.6137 | 1.6 | < 10 | 1.7 | 1.9 | 1.5 | 1.8 | 1.7 | 1.7 | 1.8 | 1.8 | 1.837 | < 1 | 1.7 | < 5 | 1.7 | 1.5951 | 2.4187 | [5 empty] | 107
| 1 | 6 | 1.76 | 2.6425 | 1.52 | < 10 | 1.7 | 1.9 | 2 | 1.7 | 1.7 | 1.7 | 1.7 | 1.8 | 1.8892 | < 1 | 1.8 | < 5 | 1.6 | 1.4544 | 2.4651
| Mean |  | 1.7783 | 2.5879 | 1.6233 |  | 1.6833 | 1.8667 | 1.9 | 1.7333 | 1.7333 | 1.7 | 1.75 | 1.8 | 1.8694 |  | 1.8 |  | 1.7167 | 1.5082 | 2.6158
| Median |  | 1.775 | 2.5939 | 1.62 |  | 1.7 | 1.9 | 1.95 | 1.7 | 1.7 | 1.7 | 1.75 | 1.8 | 1.8733 |  | 1.8 |  | 1.7 | 1.5248 | 2.6098
| Std Dev. |  | 0.037639 | 0.042346 | 0.065929 |  | 0.040825 | 0.05164 | 0.20976 | 0.05164 | 0.05164 | 0 | 0.054772 | 0 | 0.020277 |  | 0.063246 |  | 0.075277 | 0.13234 | 0.17673
| Rel.Std.Dev. |  | 2.117% | 1.636% | 4.061% |  | 2.425% | 2.766% | 11.04% | 2.979% | 2.979% | 0% | 3.13% | 0% | 1.085% |  | 3.514% |  | 4.385% | 8.774% | 6.756%
| PDM3 |  | 1.281% | 47.39% | -7.546% |  | -4.129% | 6.312% | 8.211% | -1.282% | -1.282% | -3.18% | -0.3324% | 2.515% | 6.468% |  | 2.515% |  | -2.231% | -14.1% | 48.98%
| Z-Score (Absolute) |  | 0 | 5.99 | 1.15 | 23.86 | 0.7 | 0.65 | 0.9 | 0.33 | 0.33 | 0.58 | 0.21 | 0.16 | 0.67 | 9.47 | 0.16 | 5.34 | 0.46 | 2 | 6.2
| Analytical results for Sr in OREAS 235b (Certified Value 40.5 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 40.57 | 42.6308 | 50.7 | 54.58 | 42.8 | 44 | 51.7 | 39 | 39 | 38 | 40.2 | 42.1 | 40.3 | 42.4962 | 38 | 32.81 | 39 | 36.9 | 44.5 | 27.1 | 122.4892 | 41.8026 |  |  | 1
| 1 | 2 | 39.75 | 42.1145 | 51.4 | 54.68 | 42.9 | 38 | 50.8 | 40 | 41 | 38.6 | 38.9 | 42 | 42.8 | 42.9873 | 37 | 31.68 | 41.5 | 35.4 | 46.6 | 26.6 | 129.7426 | 39.5321 |  |  | 27
| 1 | 3 | 41.27 | 42.115 | 50.4 | 55.2133 | 42 | 39 | 52.8 | 42 | 41.1 | 36.9 | 39.7 | 41.9 | 42.7 | 43.5078 | 38 | 37.8 | 39.2 | 34.4 | 41.8 | 26.4 | 129.5698 | 41.6677 |  |  | 16
| 1 | 4 | 41.35 | 42.7212 | 53.9 | 54.0933 | 42.2 | 41 | 50.3 | 39 | 41.4 | 37 | 40.8 | 43.4 | 40.6 | 41.7445 | 38 | 33.77 | 39.4 | 37.8 | 41 | 26.4 | 129.9548 | 41.4177 |  |  | 40.5041
| 1 | 5 | 40.04 | 41.1135 | 51.4 | 54.16 | 42.5 | 42 | 50.2 | 39 | 40.4 | 38.3 | 41.3 | 43.4 | 39.7 | 41.1253 | 40 | 32.94 | 36.7 | 37.5 | 42.7 | 26.8 | 129.401 | 38.5446 |  |  | 108
| 1 | 6 | 40.39 | 42.3905 | 52.3 | 54.8433 | 42.6 | 42 | 52.1 | 40 | 40 | 38 | 40.2 | 41.9 | 42.7 | 42.8036 | 39 | 32.66 | 38.1 | 35.9 | 42.4 | 27.4 | 129.5412 | 39.207
| Mean |  | 40.5617 | 42.1809 | 51.6833 | 54.595 | 42.5 | 41 | 51.3167 | 39.8333 | 40.4833 | 37.8 | 40.1833 | 42.45 | 41.4667 | 42.4441 | 38.3333 | 33.61 | 38.9833 | 36.3167 | 43.1667 | 26.7833 | 128.4498 | 40.3619
| Median |  | 40.48 | 42.2528 | 51.4 | 54.63 | 42.55 | 41.5 | 51.25 | 39.5 | 40.7 | 38 | 40.2 | 42.05 | 41.65 | 42.6499 | 38 | 32.875 | 39.1 | 36.4 | 42.55 | 26.7 | 129.5555 | 40.4749
| Std Dev. |  | 0.64555 | 0.58077 | 1.2703 | 0.42246 | 0.34641 | 2.1909 | 1.0496 | 1.169 | 0.88638 | 0.6957 | 0.83766 | 0.73959 | 1.418 | 0.87032 | 1.0328 | 2.1586 | 1.5842 | 1.3136 | 2.0461 | 0.40208 | 2.9263 | 1.4297
| Rel.Std.Dev. |  | 1.592% | 1.377% | 2.458% | 0.7738% | 0.8151% | 5.344% | 2.045% | 2.935% | 2.189% | 1.84% | 2.085% | 1.742% | 3.42% | 2.051% | 2.694% | 6.423% | 4.064% | 3.617% | 4.74% | 1.501% | 2.278% | 3.542%
| PDM3 |  | 0.1422% | 4.14% | 27.6% | 34.79% | 4.928% | 1.224% | 26.7% | -1.656% | -0.05122% | -6.676% | -0.7919% | 4.804% | 2.377% | 4.79% | -5.359% | -17.02% | -3.755% | -10.34% | 6.574% | -33.87% | 217.1% | -0.351%
| Z-Score (Absolute) |  | 0.08 | 0.54 | 4.18 | 5.3 | 0.66 | 0.08 | 4.04 | 0.36 | 0.11 | 1.14 | 0.23 | 0.64 | 0.26 | 0.64 | 0.94 | 2.75 | 0.69 | 1.71 | 0.91 | 5.37 | 33.62 | 0.16
| Analytical results for Ta in OREAS 235b (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.01 | < 0.05 | < 5 | < 0.05 | 0.06 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0518 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0326 | [10 empty] | 28
| 1 | 3 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.07 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0614 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.06 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0511 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.3437 | [10 empty] | 109
| 1 | 6 | 0.01 | < 0.05 | < 5 | < 0.05 | 0.06 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0533
| Mean |  | 1.01: 0.01; 1.09: 0.061667; 1.12: 0.01; 1.13: 0.01; 1.19: 0.098983
| Median |  | 1.01: 0.01; 1.09: 0.06; 1.12: 0.01; 1.13: 0.01; 1.19: 0.05255
| Std Dev. |  | 1.01: 0; 1.09: 0.0075277; 1.12: 0; 1.13: 0; 1.19: 0.12026
| Rel.Std.Dev. |  | 1.01: 0%; 1.09: 12.21%; 1.12: 0%; 1.13: 0%; 1.19: 121.5%
PDM3
| Z-Score (Absolute) |  | 0.64 | 0 | 87.07 | 0 | 1.29 | 0 | 0.88 | 0.53 | 0.53 | 0.7 | 0.7 | 0 | 0.7 | 2.6
| Analytical results for Tb in OREAS 235b (Indicative Value 0.51 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.21; 1.09: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*OES |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.51432; 1.06: 0.4; 1.07: 0.483; 1.08: 0.47; 1.09: 1.64 |  |  |  | 1
| 1 | 2 | 1.01: 0.51691; 1.06: 0.4; 1.07: 0.464; 1.08: 0.47; 1.09: 1.6 |  |  |  | 11
| 1 | 3 | 1.01: 0.52431; 1.06: 0.4; 1.07: 0.512; 1.08: 0.47; 1.09: 1.56 |  |  |  | 16
| 1 | 4 | 1.01: 0.51446; 1.06: 0.4; 1.07: 0.484; 1.08: 0.55; 1.09: 1.56 |  |  |  | 0.50544
| 1 | 5 | 1.01: 0.52763; 1.06: 0.4; 1.07: 0.47; 1.08: 0.55; 1.09: 1.58 |  |  |  | 17
| 1 | 6 | 1.01: 0.52922; 1.06: 0.4; 1.07: 0.508; 1.08: 0.54; 1.09: 1.62
| Mean |  | 1.01: 0.52114; 1.06: 0.4; 1.07: 0.48683; 1.08: 0.50833; 1.09: 1.5933
| Median |  | 1.01: 0.52061; 1.06: 0.4; 1.07: 0.4835; 1.08: 0.505; 1.09: 1.59
| Std Dev. |  | 1.01: 0.0067327; 1.06: 6.0809e-17; 1.07: 0.019539; 1.08: 0.042151; 1.09: 0.03266
| Rel.Std.Dev. |  | 1.01: 1.292%; 1.06: 1.52e-14%; 1.07: 4.013%; 1.08: 8.292%; 1.09: 2.05%
| PDM3 |  | 1.01: 3.107%; 1.06: -20.86%; 1.07: -3.68%; 1.08: 0.5734%; 1.09: 215.2%
| Z-Score (Absolute) |  | 1.01: 0.4; 1.06: 3.4; 1.07: 0.67; 1.08: 0; 1.09: 34.03
| Analytical results for Te in OREAS 235b (Indicative Value 0.029 ppm) | [25 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | < 0.02 | 0.05 | 5.6033 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.01 | 0.03 | < 0.05 | 0.04 | [10 empty] | 1
| 1 | 2 | < 0.02 | 0.04 | 5.434 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.02 | < 0.05 | 0.05 | 0.0239 | [9 empty] | 6
| 1 | 3 | < 0.02 | 0.03 | 6.8152 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | 0.02 | < 0.05 | 0.06 | 0.0239 | [9 empty] | 16
| 1 | 4 | < 0.02 | 0.07 | 6.6347 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.02 | < 0.05 | 0.04 | 0.0717 | [9 empty] | 0.028952
| 1 | 5 | < 0.02 | 0.04 | 5.8345 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | < 0.02 | < 0.05 | 0.05 | [10 empty] | 18
| 1 | 6 | < 0.02 | 0.06 | 6.5397 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.01 | < 0.02 | < 0.05 | 0.05 | 0.0239
| Mean |  |  | 0.048333 | 6.1436 |  | 0.04 |  |  | 0.02 | 0.02 | 0.021667 | 0.013333 | 0.025 |  | 0.048333 | 0.03585
| Median |  |  | 0.045 | 6.1871 |  | 0.04 |  |  | 0.02 | 0.02 | 0.02 | 0.01 | 0.025 |  | 0.05 | 0.0239
| Std Dev. |  |  | 0.01472 | 0.58991 |  | 0 |  |  | 0 | 0 | 0.0040825 | 0.005164 | 0.0070711 |  | 0.0075277 | 0.0239
| Rel.Std.Dev. |  |  | 30.45% | 9.602% |  | 0% |  |  | 0% | 0% | 18.84% | 38.73% | 28.28% |  | 15.57% | 66.67%
| PDM3 |  |  | 66.94% | 21120% |  | 38.16% |  |  | -30.92% | -30.92% | -25.16% | -53.95% | -13.65% |  | 66.94% | 23.83%
| Z-Score (Absolute) |  | 0.93 | 1.45 | 380.26 | 0 | 0.93 | 0 | 1.55 | 0.31 | 0.31 | 0.21 | 0.73 | 0.62 | 0 | 1.45 | 0.67
| Analytical results for Th in OREAS 235b (Certified Value 12.9 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 12.332 | 13.0463 | 13.5 | 13.54 | 13.7 | 15.6 | 13.7 | 12.6 | 12.6 | 12.7 | 12.4 | 11.8 | 12.6491 | 12.75 | 12.8 | 19.98 | 15.9 | 13.3 | 8.9855 | 10.0167 |  |  |  |  | 1
| 1 | 2 | 12.008 | 12.6218 | 13.4 | 13.71 | 12.4 | 15.8 | 13.6 | 12.9 | 12.5 | 12.4 | 12.6 | 11.8 | 13.4164 | 12.27 | 13 | 18.74 | 16.5 | 14 | 8.1338 | 11.2486 |  |  |  |  | 31
| 1 | 3 | 12.413 | 12.4172 | 13 | 13.47 | 11.6 | 15.8 | 13.4 | 12.8 | 12.4 | 12.7 | 12.8 | 12 | 13.3453 | 12.81 | 13.2 | 18.45 | 16.4 | 12.9 | 10.5141 | 11.1325 |  |  |  |  | 16
| 1 | 4 | 12.497 | 12.5386 | 13 | 13.63 | 12.7 | 15.5 | 13.2 | 13.1 | 12.6 | 12.8 | 12.5 | 12.1 | 12.8144 | 12.54 | 13.6 | 18.3 | 16.4 | 12.7 | 10.3543 | 11.3241 |  |  |  |  | 12.8583
| 1 | 5 | 12.108 | 12.5207 | 12.8 | 13.53 | 13.4 | 15.8 | 13.2 | 13 | 12.7 | 13.2 | 12.7 | 11.9 | 13.6287 | 12.26 | 13.6 | 14.87 | 16.6 | 13.2 | 10.5608 | 11.0467 |  |  |  |  | 110
| 1 | 6 | 12.359 | 12.4591 | 12.7 | 13.67 | 13.3 | 15.8 | 13.4 | 12.7 | 12.8 | 12.8 | 12.9 | 12.7 | 14.3204 | 13.19 | 12.9 | 16.88 | 16.3 | 13 | 8.8587 | 11.341
| Mean |  | 12.2862 | 12.6006 | 13.0667 | 13.5917 | 12.85 | 15.7167 | 13.4167 | 12.85 | 12.6 | 12.7667 | 12.65 | 12.05 | 13.3624 | 12.6367 | 13.1833 | 17.87 | 16.35 | 13.1833 | 9.5679 | 11.0183
| Median |  | 12.3455 | 12.5297 | 13 | 13.585 | 13 | 15.8 | 13.4 | 12.85 | 12.6 | 12.75 | 12.65 | 11.95 | 13.3809 | 12.645 | 13.1 | 18.375 | 16.4 | 13.1 | 9.6699 | 11.1905
| Std Dev. |  | 0.18818 | 0.22933 | 0.32042 | 0.09261 | 0.77653 | 0.13292 | 0.20412 | 0.18708 | 0.14142 | 0.2582 | 0.18708 | 0.33912 | 0.60015 | 0.3563 | 0.34881 | 1.7726 | 0.2429 | 0.4535 | 1.0391 | 0.50354
| Rel.Std.Dev. |  | 1.532% | 1.82% | 2.452% | 0.6814% | 6.043% | 0.8457% | 1.521% | 1.456% | 1.122% | 2.022% | 1.479% | 2.814% | 4.491% | 2.82% | 2.646% | 9.919% | 1.486% | 3.44% | 10.86% | 4.57%
| PDM3 |  | -4.45% | -2.004% | 1.62% | 5.703% | -0.06483% | 22.23% | 4.342% | -0.06483% | -2.009% | -0.7129% | -1.62% | -6.286% | 3.92% | -1.724% | 2.528% | 38.98% | 27.15% | 2.528% | -25.59% | -14.31%
| Z-Score (Absolute) |  | 0.9 | 0.4 | 0.35 | 1.18 | 0 | 4.57 | 0.9 | 0 | 0.4 | 0.13 | 0.32 | 1.28 | 0.82 | 0.34 | 0.53 | 8.01 | 5.58 | 0.53 | 5.23 | 2.92
| Analytical results for Ti in OREAS 235b (Certified Value 0.193 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1807 | 0.18 | 0.16767 | 0.2121 | 0.203 | 0.239 | 0.184 | 0.194 | 0.216 | 0.198 | 0.191 | 0.187 | 0.2168 | 0.192 | 0.19 | 0.18676 | 0.3 | 0.214 | 0.18885 | 0.3175 | 0.19806 |  |  |  | 1
| 1 | 2 | 0.1796 | 0.18 | 0.16767 | 0.2091 | 0.197 | 0.238 | 0.186 | 0.199 | 0.213 | 0.201 | 0.191 | 0.192 | 0.21621 | 0.192 | 0.19 | 0.18543 | 0.29 | 0.213 | 0.18511 | 0.3197 | 0.19418 |  |  |  | 32
| 1 | 3 | 0.1841 | 0.17 | 0.174 | 0.2041 | 0.189 | 0.239 | 0.183 | 0.197 | 0.212 | 0.201 | 0.193 | 0.191 | 0.21452 | 0.198 | 0.19 | 0.18654 | 0.3 | 0.198 | 0.18591 | 0.3218 | 0.1993 |  |  |  | 16
| 1 | 4 | 0.1794 | 0.17 | 0.16433 | 0.206 | 0.195 | 0.241 | 0.181 | 0.198 | 0.212 | 0.201 | 0.19 | 0.187 | 0.21686 | 0.192 | 0.19 | 0.18663 | 0.29 | 0.202 | 0.18488 | 0.3178 | 0.20359 |  |  |  | 0.19283
| 1 | 5 | 0.1773 | 0.16 | 0.16767 | 0.206 | 0.199 | 0.246 | 0.182 | 0.198 | 0.209 | 0.197 | 0.191 | 0.183 | 0.2183 | 0.198 | 0.17 | 0.18867 | 0.3 | 0.205 | 0.1863 | 0.2971 | 0.19673 |  |  |  | 111
| 1 | 6 | 0.1779 | 0.16 | 0.17067 | 0.2103 | 0.199 | 0.239 | 0.186 | 0.195 | 0.211 | 0.198 | 0.196 | 0.192 | 0.21611 | 0.198 | 0.18 | 0.19034 | 0.3 | 0.21 | 0.18757 | 0.3269 | 0.19473
| Mean |  | 0.17983 | 0.17 | 0.16867 | 0.20793 | 0.197 | 0.24033 | 0.18367 | 0.19683 | 0.21217 | 0.19933 | 0.192 | 0.18867 | 0.21647 | 0.195 | 0.185 | 0.18739 | 0.29667 | 0.207 | 0.18644 | 0.3168 | 0.19777
| Median |  | 0.1795 | 0.17 | 0.16767 | 0.20755 | 0.198 | 0.239 | 0.1835 | 0.1975 | 0.212 | 0.1995 | 0.191 | 0.189 | 0.21651 | 0.195 | 0.19 | 0.18669 | 0.3 | 0.2075 | 0.1861 | 0.31875 | 0.1974
| Std Dev. |  | 0.0024229 | 0.0089443 | 0.0032931 | 0.0030494 | 0.0047329 | 0.0029439 | 0.0020656 | 0.0019408 | 0.0023166 | 0.0018619 | 0.0021909 | 0.0036148 | 0.0012318 | 0.0032863 | 0.0083666 | 0.0017803 | 0.005164 | 0.0063875 | 0.0015224 | 0.010247 | 0.0034498
| Rel.Std.Dev. |  | 1.347% | 5.261% | 1.952% | 1.467% | 2.402% | 1.225% | 1.125% | 0.986% | 1.092% | 0.9341% | 1.141% | 1.916% | 0.569% | 1.685% | 4.522% | 0.95% | 1.741% | 3.086% | 0.8166% | 3.235% | 1.744%
| PDM3 |  | -6.741% | -11.84% | -12.53% | 7.831% | 2.161% | 24.63% | -4.754% | 2.074% | 10.03% | 3.371% | -0.432% | -2.161% | 12.26% | 1.124% | -4.062% | -2.821% | 53.85% | 7.347% | -3.317% | 64.29% | 2.558%
| Z-Score (Absolute) |  | 1.03 | 1.63 | 1.71 | 0.67 | 0.01 | 2.64 | 0.8 | 0 | 0.93 | 0.15 | 0.29 | 0.5 | 1.19 | 0.12 | 0.72 | 0.57 | 6.06 | 0.62 | 0.63 | 7.29 | 0.06
| Analytical results for Tl in OREAS 235b (Certified Value 0.53 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.52 | 0.55 | < 10 | 0.58 | 0.5 | 0.59 | 0.57 | 0.5 | 0.52 | 0.5 | 0.55 | 0.51 | 0.5831 | 0.761 | 0.49 | < 5 | 0.52 | 0.4418 | 0.6048 | [5 empty] | 1
| 1 | 2 | 0.52 | 0.55 | < 10 | 0.57 | 0.37 | 0.59 | 0.56 | 0.5 | 0.52 | 0.49 | 0.56 | 0.52 | 0.59616 | 0.77 | 0.49 | < 5 | 0.49 | 0.4068 | 0.6407 | [5 empty] | 33
| 1 | 3 | 0.52 | 0.53 | < 10 | 0.57 | 0.4 | 0.56 | 0.56 | 0.52 | 0.53 | 0.48 | 0.56 | 0.51 | 0.59448 | 0.755 | 0.51 | < 5 | 0.52 | 0.5002 | 0.6508 | [5 empty] | 16
| 1 | 4 | 0.55 | 0.53 | < 10 | 0.58 | 0.43 | 0.56 | 0.55 | 0.52 | 0.53 | 0.49 | 0.56 | 0.54 | 0.58402 | 0.767 | 0.52 | < 5 | 0.5 | 0.5038 | 0.6197 | [5 empty] | 0.53182
| 1 | 5 | 0.51 | 0.53 | < 10 | 0.57 | 0.45 | 0.61 | 0.55 | 0.51 | 0.5 | 0.53 | 0.56 | 0.53 | 0.56797 | 0.764 | 0.49 | < 5 | 0.49 | 0.4973 | 0.6206 | [5 empty] | 112
| 1 | 6 | 0.52 | 0.53 | < 10 | 0.58 | 0.44 | 0.55 | 0.55 | 0.49 | 0.53 | 0.5 | 0.55 | 0.54 | 0.59021 | 0.808 | 0.55 | < 5 | 0.52 | 0.433 | 0.6371
| Mean |  | 0.52333 | 0.53667 |  | 0.575 | 0.43167 | 0.57667 | 0.55667 | 0.50667 | 0.52167 | 0.49833 | 0.55667 | 0.525 | 0.58599 | 0.77083 | 0.50833 |  | 0.50667 | 0.46382 | 0.62895
| Median |  | 0.52 | 0.53 |  | 0.575 | 0.435 | 0.575 | 0.555 | 0.505 | 0.525 | 0.495 | 0.56 | 0.525 | 0.58712 | 0.7655 | 0.5 |  | 0.51 | 0.46955 | 0.62885
| Std Dev. |  | 0.013663 | 0.010328 |  | 0.0054772 | 0.04446 | 0.023381 | 0.008165 | 0.012111 | 0.01169 | 0.017224 | 0.005164 | 0.013784 | 0.010298 | 0.018925 | 0.024014 |  | 0.015055 | 0.041782 | 0.016855
| Rel.Std.Dev. |  | 2.611% | 1.924% |  | 0.9526% | 10.3% | 4.054% | 1.467% | 2.39% | 2.241% | 3.456% | 0.9277% | 2.626% | 1.757% | 2.455% | 4.724% |  | 2.971% | 9.008% | 2.68%
| PDM3 |  | -1.595% | 0.9117% |  | 8.12% | -18.83% | 8.433% | 4.672% | -4.729% | -1.909% | -6.296% | 4.672% | -1.282% | 10.19% | 44.94% | -4.416% |  | -4.729% | -12.79% | 18.26%
| Z-Score (Absolute) |  | 0.23 | 0 | 78.51 | 0.67 | 1.85 | 0.7 | 0.35 | 0.53 | 0.26 | 0.67 | 0.35 | 0.21 | 0.87 | 4.12 | 0.5 | 34.54 | 0.53 | 1.28 | 1.62
| Analytical results for Tm in OREAS 235b (Indicative Value 0.12 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | [9 empty] | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 | [9 empty] | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS | [9 empty] | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g | [9 empty] | 2
| 1 | 1 | 1.01: 0.14999; 1.06: 0.1; 1.07: 0.128; 1.08: 0.1159 | [9 empty] | 1
| 1 | 2 | 1.01: 0.14315; 1.06: 0.1; 1.07: 0.122; 1.08: 0.1071 | [9 empty] | 13
| 1 | 3 | 1.01: 0.13964; 1.06: 0.1; 1.07: 0.136; 1.08: 0.134 | [9 empty] | 16
| 1 | 4 | 1.01: 0.13786; 1.06: 0.1; 1.07: 0.129; 1.08: 0.1313 | [9 empty] | 0.12344
| 1 | 5 | 1.01: 0.13631; 1.06: 0.1; 1.07: 0.125; 1.08: 0.1353 | [9 empty] | 19
| 1 | 6 | 1.01: 0.1406; 1.06: 0.1; 1.07: 0.135; 1.08: 0.1163
| Mean |  | 1.01: 0.14126; 1.06: 0.1; 1.07: 0.12917; 1.08: 0.12332
| Median |  | 1.01: 0.14012; 1.06: 0.1; 1.07: 0.1285; 1.08: 0.1238
| Std Dev. |  | 1.01: 0.004878; 1.06: 1.5202e-17; 1.07: 0.0054924; 1.08: 0.011736
| Rel.Std.Dev. |  | 1.01: 3.453%; 1.06: 1.52e-14%; 1.07: 4.252%; 1.08: 9.517%
| PDM3 |  | 1.01: 14.44%; 1.06: -18.99%; 1.07: 4.643%; 1.08: -0.09607%
| Z-Score (Absolute) |  | 1.01: 1.13; 1.06: 1.97; 1.07: 0.22; 1.08: 0.22
| Analytical results for U in OREAS 235b (Certified Value 1.52 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.426 | 1.6775 | 1.6 | 1.7 | 1.54 | 2.1 | 1.6 | 1.37 | 1.46 | 1.38 | 1.55 | 1.39 | 1.5523 | 1 | 1.59 | 1.5 | 1.68 | 1.2399 | 3.3935 | [5 empty] | 1
| 1 | 2 | 1.418 | 1.6403 | 1.6 | 1.72 | 1.37 | 2.05 | 1.6 | 1.4 | 1.45 | 1.37 | 1.59 | 1.54 | 1.5457 | 1 | 1.516 | 1.59 | 1.57 | 1.1495 | 3.5114 | [5 empty] | 35
| 1 | 3 | 1.433 | 1.589 | 1.5 | 1.67 | 1.44 | 2.09 | 1.6 | 1.41 | 1.42 | 1.38 | 1.59 | 1.47 | 1.6196 | 1 | 1.719 | 1.55 | 1.61 | 1.4078 | 3.4936 | [5 empty] | 16
| 1 | 4 | 1.435 | 1.5201 | 1.5 | 1.68 | 1.45 | 2.01 | 1.5 | 1.43 | 1.46 | 1.39 | 1.55 | 1.51 | 1.6149 | 1 | 1.589 | 1.59 | 1.59 | 1.3976 | 3.5317 | [5 empty] | 1.5164
| 1 | 5 | 1.416 | 1.6569 | 1.5 | 1.68 | 1.53 | 2.05 | 1.5 | 1.43 | 1.5 | 1.44 | 1.59 | 1.44 | 1.6648 | 1 | 1.569 | 1.43 | 1.57 | 1.4164 | 3.5419 | [5 empty] | 113
| 1 | 6 | 1.433 | 1.5779 | 1.5 | 1.7 | 1.51 | 2.07 | 1.6 | 1.38 | 1.46 | 1.42 | 1.59 | 1.51 | 1.541 | 1 | 1.618 | 1.49 | 1.67 | 1.2507 | 3.4694
| Mean |  | 1.4268 | 1.6103 | 1.5333 | 1.6917 | 1.4733 | 2.0617 | 1.5667 | 1.4033 | 1.4583 | 1.3967 | 1.5767 | 1.4767 | 1.5897 | 1 | 1.6002 | 1.525 | 1.615 | 1.3103 | 3.4903
| Median |  | 1.4295 | 1.6147 | 1.5 | 1.69 | 1.48 | 2.06 | 1.6 | 1.405 | 1.46 | 1.385 | 1.59 | 1.49 | 1.5836 | 1 | 1.5895 | 1.525 | 1.6 | 1.3241 | 3.5025
| Std Dev. |  | 0.0082321 | 0.058689 | 0.05164 | 0.018348 | 0.06532 | 0.032506 | 0.05164 | 0.025033 | 0.025626 | 0.027325 | 0.020656 | 0.055015 | 0.050758 | 0 | 0.067413 | 0.063166 | 0.048888 | 0.11203 | 0.05411
| Rel.Std.Dev. |  | 0.5769% | 3.645% | 3.368% | 1.085% | 4.433% | 1.577% | 3.296% | 1.784% | 1.757% | 1.956% | 1.31% | 3.726% | 3.193% | 0% | 4.213% | 4.142% | 3.027% | 8.55% | 1.55%
| PDM3 |  | -5.907% | 6.193% | 1.117% | 11.56% | -2.84% | 35.96% | 3.315% | -7.456% | -3.829% | -7.896% | 3.974% | -2.62% | 4.835% | -34.05% | 5.524% | 0.5671% | 6.502% | -13.59% | 130.2%
| Z-Score (Absolute) |  | 1.11 | 0.54 | 0.15 | 1.27 | 0.69 | 4.6 | 0.15 | 1.32 | 0.82 | 1.38 | 0.24 | 0.66 | 0.36 | NA | 0.45 | 0.22 | 0.58 | 2.15 | 17.44
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 235b (Certified Value 81 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 84.6 | 86.35 | 68 | 77.3633 | 81 | 84 | 82 | 91 | 78 | 83 | 81 | 77 | 81 | 81.7814 | 85 | 80 | 79.83 | 80 | 77 | 76.1 | 92.5476 | 83.1106 |  |  | 1
| 1 | 2 | 83 | 85.58 | 69 | 77.5133 | 78 | 81 | 82 | 90 | 79 | 81 | 81 | 76 | 82 | 83.3372 | 84 | 82 | 78.29 | 79 | 77 | 74 | 97.1302 | 83.1227 |  |  | 36
| 1 | 3 | 84.6 | 85.33 | 66 | 78.0933 | 76 | 81 | 82 | 89 | 78 | 82 | 82 | 77 | 81 | 84.4613 | 86 | 81 | 83.65 | 79.6 | 75 | 73.9 | 99.0695 | 82.9263 |  |  | 16
| 1 | 4 | 81.6 | 86.33 | 69 | 77.4933 | 76 | 83 | 82 | 90 | 79 | 80 | 81 | 77 | 81 | 79.7864 | 83 | 81 | 80.76 | 79.2 | 74 | 73.2 | 98.6863 | 82.9655 |  |  | 80.8636
| 1 | 5 | 84.7 | 84.02 | 66 | 77.3167 | 77 | 82 | 85 | 91 | 79 | 81 | 81 | 78 | 79 | 78.0636 | 85 | 78 | 79.71 | 78.5 | 74 | 75 | 98.3031 | 82.8212 |  |  | 114
| 1 | 6 | 83.3 | 84.09 | 65 | 78.4833 | 81 | 84 | 82 | 90 | 79 | 81 | 81 | 77 | 83 | 79.7596 | 84 | 79 | 78.33 | 79.7 | 76 | 76.1 | 99.9572 | 82.866
| Mean |  | 83.6333 | 85.2833 | 67.1667 | 77.7106 | 78.1667 | 82.5 | 82.5 | 90.1667 | 78.6667 | 81.3333 | 81.1667 | 77 | 81.1667 | 81.1982 | 84.5 | 80.1667 | 80.095 | 79.3333 | 75.5 | 74.7167 | 97.6157 | 82.9687
| Median |  | 83.95 | 85.455 | 67 | 77.5033 | 77.5 | 82.5 | 82 | 90 | 79 | 81 | 81 | 77 | 81 | 80.7839 | 84.5 | 80.5 | 79.77 | 79.4 | 75.5 | 74.5 | 98.4947 | 82.9459
| Std Dev. |  | 1.2372 | 1.0338 | 1.7224 | 0.47021 | 2.3166 | 1.3784 | 1.2247 | 0.75277 | 0.5164 | 1.0328 | 0.40825 | 0.63246 | 1.3292 | 2.427 | 1.0488 | 1.472 | 1.9835 | 0.54283 | 1.3784 | 1.2156 | 2.6509 | 0.12487
| Rel.Std.Dev. |  | 1.479% | 1.212% | 2.564% | 0.6051% | 2.964% | 1.671% | 1.485% | 0.8349% | 0.6564% | 1.27% | 0.503% | 0.8214% | 1.638% | 2.989% | 1.241% | 1.836% | 2.476% | 0.6842% | 1.826% | 1.627% | 2.716% | 0.1505%
| PDM3 |  | 3.425% | 5.466% | -16.94% | -3.899% | -3.335% | 2.024% | 2.024% | 11.5% | -2.717% | 0.5809% | 0.3748% | -4.778% | 0.3748% | 0.4138% | 4.497% | -0.8619% | -0.9505% | -1.892% | -6.633% | -7.602% | 20.72% | 2.603%
| Z-Score (Absolute) |  | 0.67 | 1.12 | 3.8 | 0.94 | 0.81 | 0.36 | 0.36 | 2.44 | 0.68 | 0.05 | 0 | 1.13 | 0 | 0.01 | 0.91 | 0.27 | 0.29 | 0.5 | 1.54 | 1.75 | 4.47 | 0.49
| Analytical results for W in OREAS 235b (Certified Value 0.47 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.55 | 0.4 | 7.816 | 0.6 | 0.42 | 0.4 | 0.5 | 0.51 | 0.55 | 0.48 | 0.51 | 0.4 | < 1 | < 1 | 0.5 | < 5 | 0.45 | 0.4447 | [6 empty] | 1
| 1 | 2 | 0.44 | 0.3 | 8.328 | 0.6 | 0.4 | 0.4 | 0.5 | 0.52 | 0.65 | 0.56 | 0.51 | 0.38 | < 1 | < 1 | 0.5 | < 5 | 0.46 | 0.3686 | [6 empty] | 37
| 1 | 3 | 0.48 | 0.4 | 8.088 | 0.6 | 0.37 | 0.5 | 0.5 | 0.54 | 0.56 | 0.48 | 0.53 | 0.37 | < 1 | < 1 | 0.5 | < 5 | 0.47 | 0.4738 | [6 empty] | 16
| 1 | 4 | 0.46 | 0.3 | 6.208 | 0.6 | 0.33 | 0.4 | 0.5 | 0.52 | 0.58 | 0.46 | 0.51 | 0.39 | < 1 | < 1 | 0.5 | < 5 | 0.44 | 0.4836 | [6 empty] | 0.4719
| 1 | 5 | 0.52 | 0.3 | 5.24 | 0.6 | 0.39 | 0.4 | 0.5 | 0.58 | 0.54 | 0.49 | 0.5 | 0.39 | < 1 | < 1 | 0.4 | < 5 | 0.48 | 0.4771 | [6 empty] | 115
| 1 | 6 | 0.54 | 0.3 | 7.68 | 0.6 | 0.41 | 0.4 | 0.5 | 0.54 | 0.57 | 0.47 | 0.56 | 0.4 | < 1 | < 1 | 0.4 | < 5 | 0.43 | 0.398
| Mean |  | 0.49833 | 0.33333 | 7.2267 | 0.6 | 0.38667 | 0.41667 | 0.5 | 0.535 | 0.575 | 0.49 | 0.52 | 0.38833 |  |  | 0.46667 |  | 0.455 | 0.44097
| Median |  | 0.5 | 0.3 | 7.748 | 0.6 | 0.395 | 0.4 | 0.5 | 0.53 | 0.565 | 0.48 | 0.51 | 0.39 |  |  | 0.5 |  | 0.455 | 0.45925
| Std Dev. |  | 0.044907 | 0.05164 | 1.2241 | 0 | 0.03266 | 0.040825 | 0 | 0.0251 | 0.03937 | 0.035777 | 0.021909 | 0.01169 |  |  | 0.05164 |  | 0.018708 | 0.047536
| Rel.Std.Dev. |  | 9.012% | 15.49% | 16.94% | 0% | 8.447% | 9.798% | 0% | 4.692% | 6.847% | 7.301% | 4.213% | 3.01% |  |  | 11.07% |  | 4.112% | 10.78%
| PDM3 |  | 5.602% | -29.36% | 1431% | 27.15% | -18.06% | -11.7% | 5.955% | 13.37% | 21.85% | 3.836% | 10.19% | -17.71% |  |  | -1.108% |  | -3.581% | -6.554%
| Z-Score (Absolute) |  | 0.01 | NA | 113.41 | NA | 1.9 | NA | NA | 0.6 | 1.28 | 0.15 | 0.35 | 1.87 | 0.01 | 0.01 | 0.55 | 33.73 | 0.74 | 0.98
Results from laboratories 1.06, 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 235b (Certified Value 9.57 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 9.62 | 8.5128 | 8.15 | 9.16 | 10.56 | 8.31 | 10.4 | 9.6 | 9.61 | 9.89 | 9.79 | 9.84 | 10.4448 | 10 | 8.856 | 10.11 | 9.77 | 8.6165 | 17.9419 | 9.8215 |  |  |  |  | 1
| 1 | 2 | 9.38 | 8.4822 | 8.26 | 9.1267 | 10.72 | 8.26 | 10.6 | 9.5 | 10.05 | 10.4 | 9.7 | 9.61 | 10.1723 | 10 | 8.545 | 9.36 | 9.33 | 8.4835 | 10.8237 | 9.6054 |  |  |  |  | 38
| 1 | 3 | 9.73 | 8.5739 | 8.13 | 9.2533 | 10.42 | 8.59 | 10.5 | 9.4 | 9.94 | 9.97 | 9.65 | 9.94 | 10.4832 | 10 | 9.345 | 10.19 | 9.54 | 8.512 | 10.896 | 9.8194 |  |  |  |  | 16
| 1 | 4 | 9.9 | 8.6279 | 8.44 | 9.1 | 10.53 | 8.69 | 10.3 | 9.4 | 10.05 | 9.83 | 9.69 | 10.05 | 10.0527 | 9 | 9.107 | 10.02 | 9.31 | 8.4075 | 10.6835 | 9.6969 |  |  |  |  | 9.5672
| 1 | 5 | 9.6 | 8.661 | 8.17 | 9.0967 | 10.6 | 8.99 | 10.3 | 9.4 | 9.96 | 10.6 | 10.05 | 9.93 | 10.5742 | 10 | 8.708 | 10.23 | 9.14 | 8.6165 | 10.2572 | 9.3285 |  |  |  |  | 116
| 1 | 6 | 9.79 | 8.4473 | 8.43 | 9.2267 | 10.61 | 8.25 | 10.6 | 9.5 | 9.82 | 9.81 | 9.6 | 9.75 | 10.226 | 10 | 9.288 | 9.52 | 9.75 | 8.759 | 11.0061 | 9.4567
| Mean |  | 9.67 | 8.5508 | 8.2633 | 9.1606 | 10.5733 | 8.515 | 10.45 | 9.4667 | 9.905 | 10.0833 | 9.7467 | 9.8533 | 10.3255 | 9.8333 | 8.9748 | 9.905 | 9.4733 | 8.5658 | 11.9347 | 9.6214
| Median |  | 9.675 | 8.5433 | 8.215 | 9.1433 | 10.58 | 8.45 | 10.45 | 9.45 | 9.95 | 9.93 | 9.695 | 9.885 | 10.3354 | 10 | 8.9815 | 10.065 | 9.435 | 8.5642 | 10.8598 | 9.6511
| Std Dev. |  | 0.18022 | 0.084252 | 0.14024 | 0.066145 | 0.09913 | 0.29636 | 0.13784 | 0.08165 | 0.16766 | 0.33357 | 0.16133 | 0.15629 | 0.20435 | 0.40825 | 0.32331 | 0.37077 | 0.25586 | 0.12425 | 2.9544 | 0.19905
| Rel.Std.Dev. |  | 1.864% | 0.9853% | 1.697% | 0.7221% | 0.9375% | 3.48% | 1.319% | 0.8625% | 1.693% | 3.308% | 1.655% | 1.586% | 1.979% | 4.152% | 3.602% | 3.743% | 2.701% | 1.451% | 24.75% | 2.069%
| PDM3 |  | 1.075% | -10.62% | -13.63% | -4.25% | 10.52% | -11% | 9.227% | -1.051% | 3.531% | 5.395% | 1.876% | 2.991% | 7.927% | 2.782% | -6.192% | 3.531% | -0.981% | -10.47% | 24.75% | 0.5664%
| Z-Score (Absolute) |  | 0 | 1.48 | 1.86 | 0.67 | 1.2 | 1.53 | 1.03 | 0.27 | 0.31 | 0.55 | 0.1 | 0.24 | 0.87 | NA | 0.92 | 0.31 | 0.26 | 1.46 | 3 | 0.06
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 235b (Certified Value 0.95 ppm) | [25 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab | [9 empty] | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.21; 1.09: 1.30; 1.10: 1.32 | [9 empty] | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*OES; 1.10: AR*OES | [9 empty] | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.5g; 1.10: 0.5g | [9 empty] | 2
| 1 | 1 | 1.01: 0.91591; 1.06: 0.9; 1.07: 0.815; 1.08: 0.9; 1.09: 1.1; 1.10: 1.7581 | [9 empty] | 1
| 1 | 2 | 1.01: 0.89484; 1.06: 0.9; 1.07: 0.781; 1.08: 0.9; 1.09: 1.05; 1.10: 1.2278 | [9 empty] | 39
| 1 | 3 | 1.01: 0.85216; 1.06: 0.8; 1.07: 0.875; 1.08: 0.9; 1.09: 1.04; 1.10: 1.2279 | [9 empty] | 16
| 1 | 4 | 1.01: 0.8628; 1.06: 0.9; 1.07: 0.814; 1.08: 0.9; 1.09: 1.04; 1.10: 1.1691 | [9 empty] | 0.95465
| 1 | 5 | 1.01: 0.84744; 1.06: 0.9; 1.07: 0.787; 1.08: 0.9; 1.09: 1.1; 1.10: 1.1397 | [9 empty] | 117
| 1 | 6 | 1.01: 0.85974; 1.06: 0.9; 1.07: 0.861; 1.08: 0.8; 1.09: 1.08; 1.10: 1.2283
| Mean |  | 1.01: 0.87215; 1.06: 0.88333; 1.07: 0.82217; 1.08: 0.88333; 1.09: 1.0683; 1.10: 1.2918
| Median |  | 1.01: 0.86127; 1.06: 0.9; 1.07: 0.8145; 1.08: 0.9; 1.09: 1.065; 1.10: 1.2279
| Std Dev. |  | 1.01: 0.027133; 1.06: 0.040825; 1.07: 0.038338; 1.08: 0.040825; 1.09: 0.028577; 1.10: 0.23145
| Rel.Std.Dev. |  | 1.01: 3.111%; 1.06: 4.622%; 1.07: 4.663%; 1.08: 4.622%; 1.09: 2.675%; 1.10: 17.92%
| PDM3 |  | 1.01: -8.642%; 1.06: -7.47%; 1.07: -13.88%; 1.08: -7.47%; 1.09: 11.91%; 1.10: 35.32%
| Z-Score (Absolute) |  | 1.01: 0.21; 1.06: 0; 1.07: 1.14; 1.08: 0; 1.09: 3.45; 1.10: 7.61
| Analytical results for Zn in OREAS 235b (Certified Value 89 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 1
| 1 | 1 | 83.5 | 91.245 | 89 | 79.5933 | 90 | 86 | 94 | 90 | 87 | 89 | 90 | 88 | 90.7 | 94.2461 | 93 | 95.71 | 86 | 91.07 | 85.7 | 89 | 104.272 | 92 | 80.0682 | 91.1191 | 1
| 1 | 2 | 82.2 | 89.089 | 85.7 | 80.5333 | 90 | 89 | 92 | 94 | 88 | 87 | 90 | 87 | 91.9 | 91.9468 | 90 | 87.9 | 92 | 89.53 | 86.2 | 88 | 101.528 | 94 | 78.3777 | 91.1121 | 40
| 1 | 3 | 84.6 | 88.581 | 85 | 81.4267 | 91 | 89 | 93 | 88 | 88 | 89 | 91 | 88 | 89.8 | 93.2872 | 92 | 105.3 | 85 | 90.53 | 85.7 | 87 | 101.2144 | 95 | 79.4021 | 91.0178 | 16
| 1 | 4 | 84 | 90.057 | 87.3 | 83.5967 | 91 | 90 | 94 | 88 | 89 | 88 | 91 | 88 | 89.8 | 89.5398 | 92 | 94.3 | 86 | 89.29 | 87.5 | 84 | 100.7664 | 96 | 80.507 | 92.4106 | 89.4789
| 1 | 5 | 83.4 | 89.804 | 84.2 | 84.0433 | 91 | 90 | 95 | 92 | 89 | 88 | 90 | 89 | 88.2 | 92.2892 | 93 | 101.8 | 82 | 91.06 | 87.2 | 85 | 102.6368 | 98 | 77.3978 | 92.4808 | 118
| 1 | 6 | 84 | 88.759 | 84.4 | 81.3733 | 90 | 93 | 93 | 90 | 88 | 87 | 91 | 88 | 91.1 | 90.6107 | 91 | 88.02 | 82 | 89.88 | 88.6 | 85 | 105.2464 | 98 | 76.6697 | 90.7943
| Mean |  | 83.6167 | 89.5892 | 85.9333 | 81.7611 | 90.5 | 89.5 | 93.5 | 90.3333 | 88.1667 | 88 | 90.5 | 88 | 90.25 | 91.9866 | 91.8333 | 95.505 | 85.5 | 90.2267 | 86.8167 | 86.3333 | 102.6107 | 95.5 | 78.7371 | 91.4891
| Median |  | 83.75 | 89.4465 | 85.35 | 81.4 | 90.5 | 89.5 | 93.5 | 90 | 88 | 88 | 90.5 | 88 | 90.25 | 92.118 | 92 | 95.005 | 85.5 | 90.205 | 86.7 | 86 | 102.0824 | 95.5 | 78.8899 | 91.1156
| Std Dev. |  | 0.81588 | 0.99625 | 1.8737 | 1.7347 | 0.54772 | 2.2583 | 1.0488 | 2.3381 | 0.75277 | 0.89443 | 0.54772 | 0.63246 | 1.2849 | 1.7187 | 1.169 | 7.0841 | 3.6742 | 0.77208 | 1.1548 | 1.9664 | 1.8017 | 2.3452 | 1.5194 | 0.75052
| Rel.Std.Dev. |  | 0.9757% | 1.112% | 2.18% | 2.122% | 0.6052% | 2.523% | 1.122% | 2.588% | 0.8538% | 1.016% | 0.6052% | 0.7187% | 1.424% | 1.868% | 1.273% | 7.418% | 4.297% | 0.8557% | 1.33% | 2.278% | 1.756% | 2.456% | 1.93% | 0.8203%
| PDM3 |  | -6.552% | 0.1232% | -3.962% | -8.625% | 1.141% | 0.02355% | 4.494% | 0.9549% | -1.467% | -1.653% | 1.141% | -1.653% | 0.8617% | 2.803% | 2.631% | 6.735% | -4.447% | 0.8357% | -2.975% | -3.515% | 14.68% | 6.729% | -12% | 2.247%
| Z-Score (Absolute) |  | 2.12 | 0.11 | 1.34 | 2.74 | 0.2 | 0.14 | 1.21 | 0.14 | 0.59 | 0.64 | 0.2 | 0.64 | 0.12 | 0.7 | 0.65 | 1.89 | 1.48 | 0.11 | 1.04 | 1.2 | 4.28 | 1.88 | 3.76 | 0.53
| Analytical results for Zr in OREAS 235b (Certified Value 20.9 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 21.01 | 34.03 | 11 | 23.7279 | 23 | 14.7 | 26.1 | 20.2 | 20.3 | 20.7 | 21.6 | 22.2 | 24.3392 | 14 | 19.5 | 23.4 | 24.4 | 15.84 | 97.596 | 21.207 |  |  |  |  | 1
| 1 | 2 | 20.48 | 33.92 | 10.1 | 23.6552 | 24 | 13 | 27.2 | 20 | 21.2 | 21.8 | 21.5 | 22.4 | 24.7271 | 13 | 19.9 | 22.9 | 26.2 | 15.4 | 96.951 | 19.5859 |  |  |  |  | 41
| 1 | 3 | 21.05 | 33.21 | 10.7 | 24.0676 | 25.4 | 12 | 25.3 | 19.9 | 21.1 | 21 | 21.9 | 22.5 | 24.1087 | 13 | 19.2 | 22.4 | 22.1 | 15.04 | 96.9201 | 21.4407 |  |  |  |  | 16
| 1 | 4 | 20.65 | 33.45 | 9.8 | 23.6412 | 22.7 | 12.6 | 26.3 | 19.7 | 21.5 | 20.4 | 21.8 | 22.8 | 24.8413 | 15 | 19.2 | 24.1 | 23.3 | 15 | 96.8239 | 21.2966 |  |  |  |  | 20.8776
| 1 | 5 | 20.06 | 33.02 | 9.9 | 23.5932 | 25.1 | 15 | 25.9 | 19.6 | 21.1 | 21.2 | 22.3 | 22.7 | 24.6411 | 14 | 17.3 | 24.8 | 24.1 | 15.02 | 95.721 | 19.0515 |  |  |  |  | 119
| 1 | 6 | 20.54 | 33.77 | 14.2 | 23.8291 | 23.5 | 16.2 | 27.1 | 20 | 20.8 | 20.8 | 21.7 | 23 | 24.2889 | 13 | 18.1 | 24.5 | 23.6 | 15.03 | 97.0889 | 19.296
| Mean |  | 20.6317 | 33.5667 | 10.95 | 23.7524 | 23.95 | 13.9167 | 26.3167 | 19.9 | 21 | 20.9833 | 21.8 | 22.6 | 24.4911 | 13.6667 | 18.8667 | 23.6833 | 23.95 | 15.2217 | 96.8501 | 20.3129
| Median |  | 20.595 | 33.61 | 10.4 | 23.6915 | 23.75 | 13.85 | 26.2 | 19.95 | 21.1 | 20.9 | 21.75 | 22.6 | 24.4901 | 13.5 | 19.2 | 23.75 | 23.85 | 15.035 | 96.9356 | 20.3965
| Std Dev. |  | 0.36788 | 0.4051 | 1.6598 | 0.17493 | 1.1041 | 1.6278 | 0.72778 | 0.21909 | 0.40988 | 0.48339 | 0.28284 | 0.28983 | 0.28672 | 0.8165 | 0.97297 | 0.9411 | 1.3605 | 0.33873 | 0.61709 | 1.1129
| Rel.Std.Dev. |  | 1.783% | 1.207% | 15.16% | 0.7365% | 4.61% | 11.7% | 2.765% | 1.101% | 1.952% | 2.304% | 1.297% | 1.282% | 1.171% | 5.974% | 5.157% | 3.974% | 5.681% | 2.225% | 0.6372% | 5.479%
| PDM3 |  | -1.178% | 60.78% | -47.55% | 13.77% | 14.72% | -33.34% | 26.05% | -4.683% | 0.5862% | 0.5064% | 4.418% | 8.25% | 17.31% | -34.54% | -9.632% | 13.44% | 14.72% | -27.09% | 363.9% | -2.705%
| Z-Score (Absolute) |  | 0.2 | 3.23 | 2.77 | 0.62 | 0.68 | 1.99 | 1.3 | 0.4 | 0.11 | 0.11 | 0.11 | 0.32 | 0.82 | 2.05 | 0.67 | 0.61 | 0.68 | 1.64 | 20.02 | 0.29
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C420: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C438: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C456: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C474: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C492: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C510: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C927: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1053: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1071: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1089: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 235b (Indicative Value 15.03 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 15.01 | 1
| 1 | 2 | 15.04 | 16
| Mean |  | 15.025 | 16
| Median |  | 15.025 | 15.025
| Std Dev. |  | 0.021213 | 22
| Rel.Std.Dev. |  | 0.1412%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 235b (Indicative Value 1.21 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.21 | 1
| 1 | 2 | 1.21 | 17
| Mean |  | 1.21 | 16
| Median |  | 1.21 | 1.21
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 235b (Indicative Value 6.06 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.06 | 1
| 1 | 2 | 6.06 | 18
| Mean |  | 6.06 | 16
| Median |  | 6.06 | 6.06
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 235b (Indicative Value 3.44 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.44 | 1
| 1 | 2 | 3.44 | 19
| Mean |  | 3.44 | 16
| Median |  | 3.44 | 3.44
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 235b (Indicative Value 3.01 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.01 | 1
| 1 | 2 | 3.01 | 16
| Mean |  | 3.01 | 16
| Median |  | 3.01 | 3.01
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 235b (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 17
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 235b (Indicative Value 1.04 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.04 | 1
| 1 | 2 | 1.04 | 18
| Mean |  | 1.04 | 16
| Median |  | 1.04 | 1.04
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 235b (Indicative Value 0.174 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.174 | 1
| 1 | 2 | 0.174 | 19
| Mean |  | 0.174 | 16
| Median |  | 0.174 | 0.174
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 235b (Indicative Value 66.81 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 66.77 | 1
| 1 | 2 | 66.84 | 16
| Mean |  | 66.805 | 16
| Median |  | 66.805 | 66.805
| Std Dev. |  | 0.049497 | 22
| Rel.Std.Dev. |  | 0.07409%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 235b (Indicative Value 0.357 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.357 | 1
| 1 | 2 | 0.356 | 17
| Mean |  | 0.3565 | 16
| Median |  | 0.3565 | 0.3565
| Std Dev. |  | 0.00070711 | 23
| Rel.Std.Dev. |  | 0.1983%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 235b (Indicative Value 0.865 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.86 | 1
| 1 | 2 | 0.87 | 18
| Mean |  | 0.865 | 16
| Median |  | 0.865 | 0.865
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.8175%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 235b (Indicative Value 2.15 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.18 | 1
| 1 | 2 | 2.11 | 21
| Mean |  | 2.145 | 16
| Median |  | 2.145 | 2.145
| Std Dev. |  | 0.049497 | 27
| Rel.Std.Dev. |  | 2.308%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 235b (Indicative Value 0.065 wt.%) |  |  |  | Indicative
C | Round | Replicate | BV Geo |  | 1
| No | No | Perth |  | wt.%
|  |  | IRC |  | 3
3
| 1 | 1 | 0.07 |  | 1
| 1 | 2 | 0.06 |  | 23
| Mean |  | 0.065 |  | 16
| Median |  | 0.065 |  | 0.065
| Std Dev. |  | 0.0070711 |  | 29
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 235b (Indicative Value 0.134 wt.%) |  |  |  | Indicative
S | Round | Replicate | Inspectorate | BV Geo | 1
| No | No | Gandhidham | Perth | wt.%
|  |  | IRC | IRC | 3
3
| 1 | 1 | 0.13 | 0.13 | 1
| 1 | 2 | 0.14 | 0.13 | 23
| 1 | 3 | 0.14 |  | 16
| 1 | 4 | 0.14 |  | 0.13417
| 1 | 5 | 0.14 |  | 29
| 1 | 6 | 0.14
| Mean |  | 0.13833 | 0.13
| Median |  | 0.14 | 0.13
| Std Dev. |  | 0.0040825 | 0
| Rel.Std.Dev. |  | 2.951% | 0%
| PDM3 |  | 3.106% | -3.106%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 235b (Indicative Value 0.2 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 25
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 235b (Indicative Value 364 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 362 | 1
| 1 | 2 | 365 | 26
| Mean |  | 363.5 | 16
| Median |  | 363.5 | 363.5
| Std Dev. |  | 2.1213 | 32
| Rel.Std.Dev. |  | 0.5836%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 235b (Indicative Value 693 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 693 | 1
| 1 | 2 | 692 | 27
| Mean |  | 692.5 | 16
| Median |  | 692.5 | 692.5
| Std Dev. |  | 0.70711 | 33
| Rel.Std.Dev. |  | 0.1021%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 235b (Indicative Value 3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 3 | 28
| Mean |  | 3 | 16
| Median |  | 3 | 3
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 235b (Indicative Value 0.37 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.38 | 1
| 1 | 2 | 0.36 | 29
| Mean |  | 0.37 | 16
| Median |  | 0.37 | 0.37
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 235b (Indicative Value 0.2 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 5
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 235b (Indicative Value 83 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 83.1 | 1
| 1 | 2 | 82.3 | 31
| Mean |  | 82.7 | 16
| Median |  | 82.7 | 82.7
| Std Dev. |  | 0.56569 | 37
| Rel.Std.Dev. |  | 0.684%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 235b (Indicative Value 18.8 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.9 | 1
| 1 | 2 | 18.6 | 32
| Mean |  | 18.75 | 16
| Median |  | 18.75 | 18.75
| Std Dev. |  | 0.21213 | 38
| Rel.Std.Dev. |  | 1.131%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 235b (Indicative Value 152 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 152 | 1
| 1 | 2 | 151 | 33
| Mean |  | 151.5 | 16
| Median |  | 151.5 | 151.5
| Std Dev. |  | 0.70711 | 39
| Rel.Std.Dev. |  | 0.4667%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 235b (Indicative Value 9.94 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.89 | 1
| 1 | 2 | 9.99 | 34
| Mean |  | 9.94 | 16
| Median |  | 9.94 | 9.94
| Std Dev. |  | 0.070711 | 40
| Rel.Std.Dev. |  | 0.7114%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 235b (Indicative Value 31 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 32 | 1
| 1 | 2 | 30 | 35
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 235b (Indicative Value 5.72 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.72 | 1
| 1 | 2 | 5.71 | 9
| Mean |  | 5.715 | 16
| Median |  | 5.715 | 5.715
| Std Dev. |  | 0.0070711 | 42
| Rel.Std.Dev. |  | 0.1237%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 235b (Indicative Value 3.38 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.4 | 1
| 1 | 2 | 3.35 | 10
| Mean |  | 3.375 | 16
| Median |  | 3.375 | 3.375
| Std Dev. |  | 0.035355 | 43
| Rel.Std.Dev. |  | 1.048%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 235b (Indicative Value 1.41 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.4 | 1
| 1 | 2 | 1.42 | 11
| Mean |  | 1.41 | 16
| Median |  | 1.41 | 1.41
| Std Dev. |  | 0.014142 | 44
| Rel.Std.Dev. |  | 1.003%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 235b (Indicative Value 19.6 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.3 | 1
| 1 | 2 | 19.9 | 39
| Mean |  | 19.6 | 16
| Median |  | 19.6 | 19.6
| Std Dev. |  | 0.42426 | 45
| Rel.Std.Dev. |  | 2.165%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 235b (Indicative Value 6.21 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.24 | 1
| 1 | 2 | 6.17 | 12
| Mean |  | 6.205 | 16
| Median |  | 6.205 | 6.205
| Std Dev. |  | 0.049497 | 46
| Rel.Std.Dev. |  | 0.7977%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 235b (Indicative Value 1.55 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.65 | 1
| 1 | 2 | 1.45 | 6
| Mean |  | 1.55 | 16
| Median |  | 1.55 | 1.55
| Std Dev. |  | 0.14142 | 47
| Rel.Std.Dev. |  | 9.124%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 235b (Indicative Value 6.38 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.35 | 1
| 1 | 2 | 6.4 | 25
| Mean |  | 6.375 | 16
| Median |  | 6.375 | 6.375
| Std Dev. |  | 0.035355 | 31
| Rel.Std.Dev. |  | 0.5546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 235b (Indicative Value 1.19 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.21 | 1
| 1 | 2 | 1.16 | 14
| Mean |  | 1.185 | 16
| Median |  | 1.185 | 1.185
| Std Dev. |  | 0.035355 | 32
| Rel.Std.Dev. |  | 2.984%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 235b (Indicative Value 0.063 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.1 | 1
| 1 | 2 | < 0.05 | 27
| Mean |  | 0.1 | 16
| Median |  | 0.1 | 0.0625
| Std Dev. |  |  | 33
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 235b (Indicative Value 42.3 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 42.5 | 1
| 1 | 2 | 42.1 | 28
| Mean |  | 42.3 | 16
| Median |  | 42.3 | 42.3
| Std Dev. |  | 0.28284 | 34
| Rel.Std.Dev. |  | 0.6687%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 235b (Indicative Value 0.49 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.48 | 29
| Mean |  | 0.49 | 16
| Median |  | 0.49 | 0.49
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 235b (Indicative Value 0.044 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.044 | 1
| 1 | 2 | 0.044 | 30
| Mean |  | 0.044 | 16
| Median |  | 0.044 | 0.044
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 235b (Indicative Value 1.2 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.2 | 31
| Mean |  | 1.2 | 16
| Median |  | 1.2 | 1.2
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 235b (Indicative Value 17.7 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.8 | 1
| 1 | 2 | 17.6 | 32
| Mean |  | 17.7 | 16
| Median |  | 17.7 | 17.7
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 0.799%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 235b (Indicative Value 38 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 37.8 | 1
| 1 | 2 | 38.1 | 9
| Mean |  | 37.95 | 16
| Median |  | 37.95 | 37.95
| Std Dev. |  | 0.21213 | 39
| Rel.Std.Dev. |  | 0.559%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 235b (Indicative Value 71 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 72 | 1
| 1 | 2 | 70 | 34
| Mean |  | 71 | 16
| Median |  | 71 | 71
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 235b (Indicative Value 21 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21 | 1
| 1 | 2 | 21 | 35
| Mean |  | 21 | 16
| Median |  | 21 | 21
| Std Dev. |  | 0 | 41
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 235b (Indicative Value 10.2 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.1 | 1
| 1 | 2 | 10.2 | 11
| Mean |  | 10.15 | 16
| Median |  | 10.15 | 10.15
| Std Dev. |  | 0.070711 | 42
| Rel.Std.Dev. |  | 0.6967%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 235b (Indicative Value 165 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 165 | 1
| 1 | 2 | 164 | 37
| Mean |  | 164.5 | 16
| Median |  | 164.5 | 164.5
| Std Dev. |  | 0.70711 | 43
| Rel.Std.Dev. |  | 0.4299%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 235b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 235b (Indicative Value 346 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 348 | 1
| 1 | 2 | 343 | 39
| Mean |  | 345.5 | 16
| Median |  | 345.5 | 345.5
| Std Dev. |  | 3.5355 | 45
| Rel.Std.Dev. |  | 1.023%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 235b (Indicative Value 15.4 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.3 | 1
| 1 | 2 | 15.4 | 40
| Mean |  | 15.35 | 16
| Median |  | 15.35 | 15.35
| Std Dev. |  | 0.070711 | 46
| Rel.Std.Dev. |  | 0.4607%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 235b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 235b (Indicative Value 7.34 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.31 | 1
| 1 | 2 | 7.36 | 10
| Mean |  | 7.335 | 16
| Median |  | 7.335 | 7.335
| Std Dev. |  | 0.035355 | 31
| Rel.Std.Dev. |  | 0.482%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 235b (Indicative Value 3.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.8 | 1
| 1 | 2 | 3.8 | 26
| Mean |  | 3.8 | 16
| Median |  | 3.8 | 3.8
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 235b (Indicative Value 148 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 148 | 1
| 1 | 2 | 148 | 27
| Mean |  | 148 | 16
| Median |  | 148 | 148
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 235b (Indicative Value 1.34 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.31 | 1
| 1 | 2 | 1.36 | 28
| Mean |  | 1.335 | 16
| Median |  | 1.335 | 1.335
| Std Dev. |  | 0.035355 | 34
| Rel.Std.Dev. |  | 2.648%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 235b (Indicative Value 0.99 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.98 | 1
| 1 | 2 | 1 | 11
| Mean |  | 0.99 | 16
| Median |  | 0.99 | 0.99
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 235b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 6
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 235b (Indicative Value 15.7 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.6 | 1
| 1 | 2 | 15.7 | 31
| Mean |  | 15.65 | 16
| Median |  | 15.65 | 15.65
| Std Dev. |  | 0.070711 | 37
| Rel.Std.Dev. |  | 0.4518%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 235b (Indicative Value 0.519 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.522 | 1
| 1 | 2 | 0.516 | 32
| Mean |  | 0.519 | 16
| Median |  | 0.519 | 0.519
| Std Dev. |  | 0.0042426 | 38
| Rel.Std.Dev. |  | 0.8175%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 235b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 33
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 235b (Indicative Value 0.53 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.52 | 1
| 1 | 2 | 0.53 | 13
| Mean |  | 0.525 | 16
| Median |  | 0.525 | 0.525
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.347%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 235b (Indicative Value 3.23 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.21 | 1
| 1 | 2 | 3.25 | 35
| Mean |  | 3.23 | 16
| Median |  | 3.23 | 3.23
| Std Dev. |  | 0.028284 | 41
| Rel.Std.Dev. |  | 0.8757%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 235b (Indicative Value 116 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 116 | 1
| 1 | 2 | 115 | 36
| Mean |  | 115.5 | 16
| Median |  | 115.5 | 115.5
| Std Dev. |  | 0.70711 | 42
| Rel.Std.Dev. |  | 0.6122%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 235b (Indicative Value 3 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 3 | 37
| Mean |  | 3 | 16
| Median |  | 3 | 3
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 235b (Indicative Value 31.4 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.5 | 1
| 1 | 2 | 31.3 | 38
| Mean |  | 31.4 | 16
| Median |  | 31.4 | 31.4
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 0.4504%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 235b (Indicative Value 3.37 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.28 | 1
| 1 | 2 | 3.46 | 39
| Mean |  | 3.37 | 16
| Median |  | 3.37 | 3.37
| Std Dev. |  | 0.12728 | 45
| Rel.Std.Dev. |  | 3.777%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 235b (Indicative Value 100 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 40
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 235b (Indicative Value 233 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 232 | 1
| 1 | 2 | 233 | 41
| Mean |  | 232.5 | 16
| Median |  | 232.5 | 232.5
| Std Dev. |  | 0.70711 | 47
| Rel.Std.Dev. |  | 0.3041%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.