## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 238b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 3.0816 | 0.084747 | 2.9121 | 3.2511 | 2.8273 | 3.3358 | 2.75% | 5.5% | 8.25% | 2.9275 | 3.2357
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 2.9986 | 0.088051 | 2.8225 | 3.1747 | 2.7345 | 3.2628 | 2.936% | 5.873% | 8.809% | 2.8487 | 3.1485
Cyanide Leach
Au, ppm | 2.7425 | 0.16769 | 2.4071 | 3.0779 | 2.2394 | 3.2456 | 6.115% | 12.23% | 18.34% | 2.6054 | 2.8796
PhotonAssay
Au, ppm | 3.164 | 0.095902 | 2.9722 | 3.3558 | 2.8763 | 3.4517 | 3.031% | 6.062% | 9.093% | 3.0058 | 3.3222
4-Acid Digestion
Ag, ppm | 0.24462 | 0.03369 | 0.17724 | 0.312 | 0.14355 | 0.34569 | 13.77% | 27.54% | 41.32% | 0.23239 | 0.25686
Al, wt.% | 7.4723 | 0.1377 | 7.1969 | 7.7477 | 7.0592 | 7.8854 | 1.843% | 3.686% | 5.528% | 7.0987 | 7.8459
As, ppm | 676.0572 | 31.539 | 612.9793 | 739.1351 | 581.4403 | 770.6741 | 4.665% | 9.33% | 14% | 642.2544 | 709.8601
Ba, ppm | 685.46 | 21.3803 | 642.6994 | 728.2206 | 621.3191 | 749.6009 | 3.119% | 6.238% | 9.357% | 651.187 | 719.733
Be, ppm | 2.4659 | 0.22755 | 2.0108 | 2.921 | 1.7833 | 3.1486 | 9.228% | 18.46% | 27.68% | 2.3426 | 2.5892
Bi, ppm | 0.40028 | 0.025647 | 0.34898 | 0.45157 | 0.32334 | 0.47722 | 6.407% | 12.81% | 19.22% | 0.38026 | 0.42029
Ca, wt.% | 0.85542 | 0.023549 | 0.80832 | 0.90252 | 0.78477 | 0.92607 | 2.753% | 5.506% | 8.259% | 0.81265 | 0.89819
Cd, ppm | 0.097127 | 0.02827 | 0.040587 | 0.15367 | 0.012317 | 0.18194 | 29.11% | 58.21% | 87.32% | 0.092271 | 0.10198
Ce, ppm | 80.6391 | 5.6196 | 69.4 | 91.8782 | 63.7804 | 97.4978 | 6.969% | 13.94% | 20.91% | 76.6071 | 84.671
Co, ppm | 17.0384 | 0.55758 | 15.9232 | 18.1536 | 15.3657 | 18.7111 | 3.273% | 6.545% | 9.818% | 16.1865 | 17.8903
Cr, ppm | 123.9533 | 13.0856 | 97.7822 | 150.1244 | 84.6966 | 163.21 | 10.56% | 21.11% | 31.67% | 117.7556 | 130.151
Cs, ppm | 9.5152 | 0.35064 | 8.8139 | 10.2164 | 8.4632 | 10.5671 | 3.685% | 7.37% | 11.06% | 9.0394 | 9.9909
Cu, ppm | 30.7977 | 1.7388 | 27.3201 | 34.2754 | 25.5813 | 36.0142 | 5.646% | 11.29% | 16.94% | 29.2579 | 32.3376
Dy, ppm | 3.5604 | 0.62567 | 2.3091 | 4.8118 | 1.6834 | 5.4375 | 17.57% | 35.15% | 52.72% | 3.3824 | 3.7385
Er, ppm | 1.8998 | 0.39771 | 1.1043 | 2.6952 | 0.70663 | 3.0929 | 20.93% | 41.87% | 62.8% | 1.8048 | 1.9947
Eu, ppm | 1.2576 | 0.11909 | 1.0195 | 1.4958 | 0.90039 | 1.6149 | 9.469% | 18.94% | 28.41% | 1.1948 | 1.3205
Fe, wt.% | 4.1016 | 0.10062 | 3.9004 | 4.3029 | 3.7998 | 4.4035 | 2.453% | 4.906% | 7.359% | 3.8965 | 4.3067
Ga, ppm | 19.5512 | 1.2072 | 17.1368 | 21.9656 | 15.9296 | 23.1728 | 6.174% | 12.35% | 18.52% | 18.5736 | 20.5288
Gd, ppm | 5.1398 | 0.37466 | 4.3905 | 5.8891 | 4.0158 | 6.2638 | 7.289% | 14.58% | 21.87% | 4.8828 | 5.3968
Hf, ppm | 3.8854 | 0.23359 | 3.4182 | 4.3526 | 3.1846 | 4.5862 | 6.012% | 12.02% | 18.04% | 3.6911 | 4.0797
Ho, ppm | 0.65576 | 0.13352 | 0.38872 | 0.9228 | 0.2552 | 1.0563 | 20.36% | 40.72% | 61.08% | 0.62297 | 0.68855
In, ppm | 0.068319 | 0.0079508 | 0.052418 | 0.084221 | 0.044467 | 0.092172 | 11.64% | 23.28% | 34.91% | 0.064903 | 0.071735
K, wt.% | 2.6916 | 0.079818 | 2.5319 | 2.8512 | 2.4521 | 2.931 | 2.966% | 5.931% | 8.897% | 2.557 | 2.8261
La, ppm | 38.8276 | 2.9567 | 32.9142 | 44.741 | 29.9575 | 47.6977 | 7.615% | 15.23% | 22.84% | 36.8862 | 40.769
Li, ppm | 48.7761 | 3.5529 | 41.6703 | 55.882 | 38.1173 | 59.435 | 7.284% | 14.57% | 21.85% | 46.3373 | 51.2149
Lu, ppm | 0.2916 | 0.046619 | 0.19836 | 0.38483 | 0.15174 | 0.43145 | 15.99% | 31.98% | 47.96% | 0.27702 | 0.30617
Mg, wt.% | 1.7445 | 0.052611 | 1.6393 | 1.8497 | 1.5867 | 1.9024 | 3.016% | 6.032% | 9.047% | 1.6573 | 1.8318
Mn, wt.% | 0.039061 | 0.0017385 | 0.035584 | 0.042538 | 0.033846 | 0.044277 | 4.451% | 8.901% | 13.35% | 0.037108 | 0.041014
Mo, ppm | 0.90149 | 0.05555 | 0.79039 | 1.0126 | 0.73484 | 1.0681 | 6.162% | 12.32% | 18.49% | 0.85641 | 0.94656
Na, wt.% | 0.67587 | 0.017352 | 0.64116 | 0.71057 | 0.62381 | 0.72792 | 2.567% | 5.135% | 7.702% | 0.64207 | 0.70966
Nb, ppm | 14.3051 | 1.038 | 12.2291 | 16.3811 | 11.1911 | 17.4191 | 7.256% | 14.51% | 21.77% | 13.5898 | 15.0203
Nd, ppm | 34.1112 | 1.7646 | 30.582 | 37.6404 | 28.8174 | 39.405 | 5.173% | 10.35% | 15.52% | 32.4056 | 35.8167
Ni, ppm | 65.2221 | 3.0898 | 59.0425 | 71.4016 | 55.9528 | 74.4914 | 4.737% | 9.475% | 14.21% | 61.961 | 68.4832
P, wt.% | 0.065288 | 0.0026914 | 0.059905 | 0.070671 | 0.057214 | 0.073362 | 4.122% | 8.245% | 12.37% | 0.062024 | 0.068553
Pb, ppm | 23.4425 | 1.3463 | 20.7499 | 26.1352 | 19.4036 | 27.4815 | 5.743% | 11.49% | 17.23% | 22.2704 | 24.6146
Pr, ppm | 9.3058 | 0.50271 | 8.3004 | 10.3112 | 7.7977 | 10.8139 | 5.402% | 10.8% | 16.21% | 8.8405 | 9.7711
Rb, ppm | 158.6634 | 9.7998 | 139.0638 | 178.263 | 129.264 | 188.0628 | 6.176% | 12.35% | 18.53% | 150.7303 | 166.5966
Re, ppm | < 0.002 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.17477 | 0.0073813 | 0.16001 | 0.18953 | 0.15263 | 0.19692 | 4.223% | 8.447% | 12.67% | 0.16603 | 0.18351
Sb, ppm | 620.187 | 26.8564 | 566.4743 | 673.8997 | 539.6179 | 700.756 | 4.33% | 8.661% | 12.99% | 589.1776 | 651.1963
Sc, ppm | 14.2704 | 1.0296 | 12.2111 | 16.3296 | 11.1815 | 17.3593 | 7.215% | 14.43% | 21.65% | 13.5569 | 14.9839
Sm, ppm | 6.5436 | 0.47783 | 5.5879 | 7.4993 | 5.1101 | 7.9771 | 7.302% | 14.6% | 21.91% | 6.2164 | 6.8708
Sn, ppm | 3.6899 | 0.26501 | 3.1599 | 4.2199 | 2.8949 | 4.4849 | 7.182% | 14.36% | 21.55% | 3.5054 | 3.8744
Sr, ppm | 117.9561 | 3.9127 | 110.1308 | 125.7814 | 106.2182 | 129.6941 | 3.317% | 6.634% | 9.951% | 112.0583 | 123.8539
Ta, ppm | 1.0702 | 0.08474 | 0.90073 | 1.2397 | 0.81599 | 1.3244 | 7.918% | 15.84% | 23.75% | 1.0167 | 1.1237
Tb, ppm | 0.67372 | 0.032661 | 0.6084 | 0.73905 | 0.57574 | 0.77171 | 4.848% | 9.696% | 14.54% | 0.64004 | 0.70741
Th, ppm | 14.513 | 0.69104 | 13.1309 | 15.8951 | 12.4399 | 16.5861 | 4.762% | 9.523% | 14.28% | 13.7873 | 15.2386
Ti, wt.% | 0.46128 | 0.017938 | 0.4254 | 0.49715 | 0.40746 | 0.51509 | 3.889% | 7.777% | 11.67% | 0.43821 | 0.48434
Tl, ppm | 0.8481 | 0.049046 | 0.75001 | 0.94619 | 0.70096 | 0.99524 | 5.783% | 11.57% | 17.35% | 0.80569 | 0.8905
Tm, ppm | 0.26676 | 0.059292 | 0.14817 | 0.38534 | 0.088881 | 0.44463 | 22.23% | 44.45% | 66.68% | 0.25342 | 0.28009
U, ppm | 2.7363 | 0.14571 | 2.4448 | 3.0277 | 2.2991 | 3.1734 | 5.325% | 10.65% | 15.98% | 2.5994 | 2.8731
V, ppm | 107.6332 | 2.9926 | 101.648 | 113.6183 | 98.6555 | 116.6109 | 2.78% | 5.561% | 8.341% | 102.2515 | 113.0148
W, ppm | 2.7974 | 0.40085 | 1.9957 | 3.5991 | 1.5949 | 4 | 14.33% | 28.66% | 42.99% | 2.6576 | 2.9373
Y, ppm | 16.0005 | 2.5138 | 10.9729 | 21.0281 | 8.4591 | 23.5419 | 15.71% | 31.42% | 47.13% | 15.2005 | 16.8006
Yb, ppm | 1.8157 | 0.38995 | 1.0358 | 2.5956 | 0.6459 | 2.9856 | 21.48% | 42.95% | 64.43% | 1.725 | 1.9065
Zn, ppm | 99.9503 | 3.8397 | 92.2708 | 107.6297 | 88.431 | 111.4695 | 3.842% | 7.683% | 11.52% | 94.9527 | 104.9478
Zr, ppm | 132.1957 | 9.9978 | 112.2001 | 152.1913 | 102.2023 | 162.1891 | 7.563% | 15.13% | 22.69% | 125.5859 | 138.8054
Aqua Regia Digestion
Ag, ppm | 0.23976 | 0.013106 | 0.21355 | 0.26598 | 0.20045 | 0.27908 | 5.466% | 10.93% | 16.4% | 0.22778 | 0.25175
Al, wt.% | 2.6181 | 0.20656 | 2.205 | 3.0312 | 1.9984 | 3.2378 | 7.89% | 15.78% | 23.67% | 2.4872 | 2.749
As, ppm | 673.6185 | 25.4678 | 622.6829 | 724.554 | 597.2151 | 750.0218 | 3.781% | 7.561% | 11.34% | 639.9376 | 707.2994
Ba, ppm | 102.5083 | 6.2452 | 90.018 | 114.9986 | 83.7728 | 121.2438 | 6.092% | 12.18% | 18.28% | 97.3829 | 107.6337
Be, ppm | 1.169 | 0.088544 | 0.99195 | 1.3461 | 0.9034 | 1.4347 | 7.574% | 15.15% | 22.72% | 1.1106 | 1.2275
Bi, ppm | 0.37697 | 0.031609 | 0.31375 | 0.44018 | 0.28214 | 0.47179 | 8.385% | 16.77% | 25.16% | 0.35812 | 0.39581
Ca, wt.% | 0.30901 | 0.019932 | 0.26914 | 0.34887 | 0.24921 | 0.36881 | 6.45% | 12.9% | 19.35% | 0.29356 | 0.32446
Cd, ppm | 0.080003 | 0.017016 | 0.045972 | 0.11403 | 0.028956 | 0.13105 | 21.27% | 42.54% | 63.81% | 0.076003 | 0.084003
Ce, ppm | 54.6957 | 3.0581 | 48.5795 | 60.812 | 45.5213 | 63.8701 | 5.591% | 11.18% | 16.77% | 51.9609 | 57.4305
Co, ppm | 15.2363 | 0.86707 | 13.5021 | 16.9704 | 12.635 | 17.8375 | 5.691% | 11.38% | 17.07% | 14.4745 | 15.9981
Cr, ppm | 106.0053 | 4.9183 | 96.1686 | 115.8419 | 91.2503 | 120.7603 | 4.64% | 9.279% | 13.92% | 100.705 | 111.3055
Cs, ppm | 6.8488 | 0.25284 | 6.3431 | 7.3544 | 6.0902 | 7.6073 | 3.692% | 7.384% | 11.08% | 6.5063 | 7.1912
Cu, ppm | 29.5904 | 1.5137 | 26.5631 | 32.6178 | 25.0495 | 34.1314 | 5.115% | 10.23% | 15.35% | 28.1109 | 31.07
Fe, wt.% | 3.5144 | 0.13442 | 3.2455 | 3.7832 | 3.1111 | 3.9176 | 3.825% | 7.65% | 11.47% | 3.3387 | 3.6901
Ga, ppm | 7.9265 | 0.57129 | 6.7839 | 9.0691 | 6.2127 | 9.6404 | 7.207% | 14.41% | 21.62% | 7.5302 | 8.3228
Ge, ppm | 0.10729 | 0.020922 | 0.065446 | 0.14913 | 0.044524 | 0.17006 | 19.5% | 39% | 58.5% | 0.10193 | 0.11265
Hf, ppm | 0.541 | 0.079563 | 0.38187 | 0.70012 | 0.30231 | 0.77969 | 14.71% | 29.41% | 44.12% | 0.51395 | 0.56805
In, ppm | 0.029382 | 0.0020957 | 0.02519 | 0.033573 | 0.023095 | 0.035669 | 7.133% | 14.27% | 21.4% | 0.027913 | 0.030851
K, wt.% | 0.87015 | 0.037101 | 0.79595 | 0.94435 | 0.75885 | 0.98145 | 4.264% | 8.527% | 12.79% | 0.82664 | 0.91366
La, ppm | 27.0633 | 2.4198 | 22.2238 | 31.9028 | 19.804 | 34.3226 | 8.941% | 17.88% | 26.82% | 25.7101 | 28.4165
Li, ppm | 38.9377 | 2.5137 | 33.9103 | 43.965 | 31.3966 | 46.4787 | 6.456% | 12.91% | 19.37% | 36.9908 | 40.8845
Lu, ppm | 0.10947 | 0.032146 | 0.04518 | 0.17376 | 0.013034 | 0.20591 | 29.36% | 58.73% | 88.09% | 0.104 | 0.11494
Mg, wt.% | 1.4147 | 0.077048 | 1.2606 | 1.5688 | 1.1836 | 1.6458 | 5.446% | 10.89% | 16.34% | 1.344 | 1.4854
Mn, wt.% | 0.027295 | 0.0015376 | 0.02422 | 0.03037 | 0.022682 | 0.031908 | 5.633% | 11.27% | 16.9% | 0.02593 | 0.02866
Mo, ppm | 0.81068 | 0.045183 | 0.72031 | 0.90104 | 0.67513 | 0.94623 | 5.573% | 11.15% | 16.72% | 0.77014 | 0.85121
Na, wt.% | 0.082073 | 0.0076589 | 0.066755 | 0.09739 | 0.059096 | 0.10505 | 9.332% | 18.66% | 28% | 0.077969 | 0.086176
Nb, ppm | 0.39277 | 0.10935 | 0.17407 | 0.61148 | 0.064716 | 0.72083 | 27.84% | 55.68% | 83.52% | 0.37313 | 0.41241
Ni, ppm | 60.6564 | 3.7933 | 53.0698 | 68.243 | 49.2765 | 72.0363 | 6.254% | 12.51% | 18.76% | 57.6235 | 63.6892
P, wt.% | 0.057782 | 0.0019417 | 0.053899 | 0.061666 | 0.051957 | 0.063607 | 3.36% | 6.721% | 10.08% | 0.054893 | 0.060671
Pb, ppm | 14.1612 | 1.7611 | 10.6391 | 17.6833 | 8.8781 | 19.4444 | 12.44% | 24.87% | 37.31% | 13.4532 | 14.8693
Rb, ppm | 83.3596 | 2.8334 | 77.6928 | 89.0264 | 74.8594 | 91.8598 | 3.399% | 6.798% | 10.2% | 79.1916 | 87.5276
Re, ppm | < 0.001 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 0.17281 | 0.012936 | 0.14694 | 0.19868 | 0.134 | 0.21162 | 7.486% | 14.97% | 22.46% | 0.16417 | 0.18145
Sb, ppm | 468.6505 | 83.3097 | 302.0312 | 635.2698 | 218.7215 | 718.5795 | 17.78% | 35.55% | 53.33% | 445.218 | 492.083
Sc, ppm | 6.2439 | 0.46356 | 5.3167 | 7.171 | 4.8532 | 7.6345 | 7.424% | 14.85% | 22.27% | 5.9317 | 6.556
Sn, ppm | 1.446 | 0.093578 | 1.2588 | 1.6331 | 1.1652 | 1.7267 | 6.472% | 12.94% | 19.41% | 1.3737 | 1.5183
Sr, ppm | 26.0732 | 1.332 | 23.4092 | 28.7372 | 22.0772 | 30.0692 | 5.109% | 10.22% | 15.33% | 24.7695 | 27.3768
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 11.9141 | 0.72431 | 10.4655 | 13.3628 | 9.7412 | 14.0871 | 6.079% | 12.16% | 18.24% | 11.3184 | 12.5098
Ti, wt.% | 0.14283 | 0.013306 | 0.11622 | 0.16944 | 0.10291 | 0.18275 | 9.316% | 18.63% | 27.95% | 0.13569 | 0.14997
Tl, ppm | 0.51836 | 0.053403 | 0.41156 | 0.62517 | 0.35815 | 0.67857 | 10.3% | 20.6% | 30.91% | 0.49245 | 0.54428
U, ppm | 1.457 | 0.13528 | 1.1865 | 1.7276 | 1.0512 | 1.8629 | 9.285% | 18.57% | 27.85% | 1.3842 | 1.5299
V, ppm | 68.5636 | 3.1735 | 62.2166 | 74.9106 | 59.0431 | 78.0841 | 4.629% | 9.257% | 13.89% | 65.1354 | 71.9918
W, ppm | 0.52 | 0.099494 | 0.32101 | 0.71899 | 0.22152 | 0.81848 | 19.13% | 38.27% | 57.4% | 0.494 | 0.546
Y, ppm | 8.9951 | 0.69732 | 7.6004 | 10.3897 | 6.9031 | 11.087 | 7.752% | 15.5% | 23.26% | 8.5453 | 9.4448
Yb, ppm | 0.86718 | 0.1428 | 0.58158 | 1.1528 | 0.43878 | 1.2956 | 16.47% | 32.93% | 49.4% | 0.82382 | 0.91053
Zn, ppm | 86.8375 | 4.2499 | 78.3377 | 95.3372 | 74.0879 | 99.5871 | 4.894% | 9.788% | 14.68% | 82.4956 | 91.1794
Zr, ppm | 21.2593 | 2.6082 | 16.0428 | 26.4757 | 13.4346 | 29.0839 | 12.27% | 24.54% | 36.81% | 20.1963 | 22.3222
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 238b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 3.0816 | 3.0514 | 3.1117 | 3.0679 | 3.0952
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 2.9986 | 2.9615 | 3.0357 | 2.984 | 3.0132
Cyanide Leach
Au, Gold (ppm) | 2.7425 | 2.6537 | 2.8313 | 2.7303 | 2.7547
PhotonAssay
Au, Gold (ppm) | 3.164 | 3.0834 | 3.2446 | 3.1599 | 3.1681
4-Acid Digestion
Ag, Silver (ppm) | 0.24462 | 0.21764 | 0.27161 | 0.2246 | 0.26465
Al, Aluminium (wt.%) | 7.4723 | 7.3373 | 7.6072 | 7.3405 | 7.604
As, Arsenic (ppm) | 676.0572 | 653.4281 | 698.6864 | 658.2617 | 693.8527
Ba, Barium (ppm) | 685.46 | 669.1373 | 701.7826 | 673.8763 | 697.0437
Be, Beryllium (ppm) | 2.4659 | 2.3067 | 2.6252 | 2.3536 | 2.5782
Bi, Bismuth (ppm) | 0.40028 | 0.36973 | 0.43083 | 0.36582 | 0.43473
Ca, Calcium (wt.%) | 0.85542 | 0.83598 | 0.87485 | 0.83908 | 0.87175
Cd, Cadmium (ppm) | 0.097127 | 0.072069 | 0.12219 | 0.09306 | 0.10119
Ce, Cerium (ppm) | 80.6391 | 77.3871 | 83.891 | 78.2944 | 82.9837
Co, Cobalt (ppm) | 17.0384 | 16.2737 | 17.8031 | 16.497 | 17.5798
Cr, Chromium (ppm) | 123.9533 | 117.1091 | 130.7975 | 119.3781 | 128.5285
Cs, Caesium (ppm) | 9.5152 | 9.1711 | 9.8592 | 9.2016 | 9.8287
Cu, Copper (ppm) | 30.7977 | 29.1962 | 32.3993 | 29.8593 | 31.7362
Dy, Dysprosium (ppm) | 3.5604 | 3.0865 | 4.0344 | 3.4092 | 3.7117
Er, Erbium (ppm) | 1.8998 | 1.613 | 2.1865 | 1.8232 | 1.9764
Eu, Europium (ppm) | 1.2576 | 1.1243 | 1.391 | 1.1941 | 1.3212
Fe, Iron (wt.%) | 4.1016 | 4.0021 | 4.2011 | 4.0497 | 4.1535
Ga, Gallium (ppm) | 19.5512 | 18.6764 | 20.426 | 19.0803 | 20.0221
Gd, Gadolinium (ppm) | 5.1398 | 4.7869 | 5.4927 | 4.9721 | 5.3075
Hf, Hafnium (ppm) | 3.8854 | 3.6594 | 4.1114 | 3.7265 | 4.0443
Ho, Holmium (ppm) | 0.65576 | 0.5457 | 0.76582 | 0.62947 | 0.68205
In, Indium (ppm) | 0.068319 | 0.062667 | 0.073972 | 0.063013 | 0.073626
K, Potassium (wt.%) | 2.6916 | 2.6347 | 2.7484 | 2.6432 | 2.7399
La, Lanthanum (ppm) | 38.8276 | 37.0677 | 40.5876 | 37.857 | 39.7983
Li, Lithium (ppm) | 48.7761 | 46.6675 | 50.8847 | 47.5382 | 50.014
Lu, Lutetium (ppm) | 0.2916 | 0.2353 | 0.34789 | 0.2726 | 0.31059
Mg, Magnesium (wt.%) | 1.7445 | 1.7082 | 1.7809 | 1.7205 | 1.7685
Mn, Manganese (wt.%) | 0.039061 | 0.037855 | 0.040268 | 0.038506 | 0.039616
Mo, Molybdenum (ppm) | 0.90149 | 0.85299 | 0.94998 | 0.84422 | 0.95876
Na, Sodium (wt.%) | 0.67587 | 0.65786 | 0.69387 | 0.66278 | 0.68895
Nb, Niobium (ppm) | 14.3051 | 13.4749 | 15.1353 | 13.8564 | 14.7538
Nd, Neodymium (ppm) | 34.1112 | 32.6168 | 35.6055 | 32.728 | 35.4944
Ni, Nickel (ppm) | 65.2221 | 63.3166 | 67.1276 | 63.4274 | 67.0167
P, Phosphorus (wt.%) | 0.065288 | 0.063535 | 0.067041 | 0.06394 | 0.066637
Pb, Lead (ppm) | 23.4425 | 22.1963 | 24.6887 | 22.3513 | 24.5338
Pr, Praseodymium (ppm) | 9.3058 | 8.8461 | 9.7654 | 8.9155 | 9.6961
Rb, Rubidium (ppm) | 158.6634 | 151.3435 | 165.9833 | 153.5466 | 163.7802
Re, Rhenium (ppm) | < 0.002 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.17477 | 0.16732 | 0.18222 | 0.16848 | 0.18107
Sb, Antimony (ppm) | 620.187 | 602.8394 | 637.5345 | 607.2902 | 633.0838
Sc, Scandium (ppm) | 14.2704 | 13.4889 | 15.0519 | 13.8201 | 14.7207
Sm, Samarium (ppm) | 6.5436 | 6.1325 | 6.9547 | 6.2782 | 6.809
Sn, Tin (ppm) | 3.6899 | 3.5198 | 3.86 | 3.5186 | 3.8612
Sr, Strontium (ppm) | 117.9561 | 114.314 | 121.5983 | 115.6398 | 120.2725
Ta, Tantalum (ppm) | 1.0702 | 1.0104 | 1.13 | 1.0348 | 1.1056
Tb, Terbium (ppm) | 0.67372 | 0.62404 | 0.72341 | 0.62962 | 0.71783
Th, Thorium (ppm) | 14.513 | 14.0221 | 15.0039 | 14.02 | 15.006
Ti, Titanium (wt.%) | 0.46128 | 0.4457 | 0.47685 | 0.44992 | 0.47263
Tl, Thallium (ppm) | 0.8481 | 0.81084 | 0.88536 | 0.81526 | 0.88093
Tm, Thulium (ppm) | 0.26676 | 0.21883 | 0.31468 | 0.25354 | 0.27997
U, Uranium (ppm) | 2.7363 | 2.6303 | 2.8422 | 2.64 | 2.8325
V, Vanadium (ppm) | 107.6332 | 105.129 | 110.1374 | 105.6719 | 109.5944
W, Tungsten (ppm) | 2.7974 | 2.5581 | 3.0368 | 2.6553 | 2.9396
Y, Yttrium (ppm) | 16.0005 | 14.8046 | 17.1965 | 15.4648 | 16.5363
Yb, Ytterbium (ppm) | 1.8157 | 1.5677 | 2.0638 | 1.7263 | 1.9052
Zn, Zinc (ppm) | 99.9503 | 95.7506 | 104.1499 | 97.521 | 102.3795
Zr, Zirconium (ppm) | 132.1957 | 125.8339 | 138.5574 | 128.8547 | 135.5367
Aqua Regia Digestion
Ag, Silver (ppm) | 0.23976 | 0.22242 | 0.25711 | 0.22217 | 0.25736
Al, Aluminium (wt.%) | 2.6181 | 2.4982 | 2.738 | 2.5805 | 2.6557
As, Arsenic (ppm) | 673.6185 | 654.6276 | 692.6094 | 662.4518 | 684.7851
Ba, Barium (ppm) | 102.5083 | 97.2584 | 107.7582 | 98.9532 | 106.0633
Be, Beryllium (ppm) | 1.169 | 1.0924 | 1.2457 | 1.1204 | 1.2176
Bi, Bismuth (ppm) | 0.37697 | 0.34576 | 0.40817 | 0.3313 | 0.42263
Ca, Calcium (wt.%) | 0.30901 | 0.29632 | 0.3217 | 0.30136 | 0.31665
Cd, Cadmium (ppm) | 0.080003 | 0.067029 | 0.092976 | 0.076328 | 0.083678
Ce, Cerium (ppm) | 54.6957 | 50.2451 | 59.1463 | 52.6047 | 56.7867
Co, Cobalt (ppm) | 15.2363 | 14.5409 | 15.9316 | 14.661 | 15.8115
Cr, Chromium (ppm) | 106.0053 | 102.4442 | 109.5663 | 103.8115 | 108.1991
Cs, Caesium (ppm) | 6.8488 | 6.4875 | 7.21 | 6.6164 | 7.0811
Cu, Copper (ppm) | 29.5904 | 28.3496 | 30.8313 | 28.8664 | 30.3145
Fe, Iron (wt.%) | 3.5144 | 3.4169 | 3.6119 | 3.4706 | 3.5581
Ga, Gallium (ppm) | 7.9265 | 7.5319 | 8.3211 | 7.6596 | 8.1935
Ge, Germanium (ppm) | 0.10729 | 0.078988 | 0.13559 | IND | IND
Hf, Hafnium (ppm) | 0.541 | 0.46826 | 0.61373 | 0.51494 | 0.56706
In, Indium (ppm) | 0.029382 | 0.027479 | 0.031284 | 0.026936 | 0.031827
K, Potassium (wt.%) | 0.87015 | 0.84462 | 0.89568 | 0.85713 | 0.88317
La, Lanthanum (ppm) | 27.0633 | 24.8845 | 29.2421 | 26.1371 | 27.9896
Li, Lithium (ppm) | 38.9377 | 37.0046 | 40.8707 | 38.0646 | 39.8107
Lu, Lutetium (ppm) | 0.10947 | 0.04445 | 0.17449 | IND | IND
Mg, Magnesium (wt.%) | 1.4147 | 1.3681 | 1.4613 | 1.3917 | 1.4377
Mn, Manganese (wt.%) | 0.027295 | 0.026334 | 0.028256 | 0.02684 | 0.02775
Mo, Molybdenum (ppm) | 0.81068 | 0.75593 | 0.86543 | 0.76179 | 0.85957
Na, Sodium (wt.%) | 0.082073 | 0.075188 | 0.088957 | 0.078568 | 0.085577
Nb, Niobium (ppm) | 0.39277 | 0.29655 | 0.48899 | 0.35897 | 0.42658
Ni, Nickel (ppm) | 60.6564 | 58.4607 | 62.8521 | 58.9458 | 62.3669
P, Phosphorus (wt.%) | 0.057782 | 0.056315 | 0.05925 | 0.056751 | 0.058813
Pb, Lead (ppm) | 14.1612 | 12.9775 | 15.345 | 13.3148 | 15.0077
Rb, Rubidium (ppm) | 83.3596 | 79.6858 | 87.0334 | 81.1631 | 85.5561
Re, Rhenium (ppm) | < 0.001 | IND | IND | IND | IND
S, Sulphur (wt.%) | 0.17281 | 0.1644 | 0.18121 | 0.16565 | 0.17996
Sb, Antimony (ppm) | 468.6505 | 426.3388 | 510.9622 | 456.5056 | 480.7953
Sc, Scandium (ppm) | 6.2439 | 5.96 | 6.5277 | 6.0909 | 6.3968
Sn, Tin (ppm) | 1.446 | 1.3411 | 1.5508 | 1.3928 | 1.4991
Sr, Strontium (ppm) | 26.0732 | 24.8693 | 27.277 | 25.4882 | 26.6582
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Th, Thorium (ppm) | 11.9141 | 11.2591 | 12.5691 | 11.5687 | 12.2596
Ti, Titanium (wt.%) | 0.14283 | 0.13538 | 0.15028 | 0.13826 | 0.1474
Tl, Thallium (ppm) | 0.51836 | 0.48403 | 0.5527 | 0.5009 | 0.53583
U, Uranium (ppm) | 1.457 | 1.3561 | 1.5579 | 1.3987 | 1.5153
V, Vanadium (ppm) | 68.5636 | 66.4293 | 70.698 | 67.2332 | 69.8941
W, Tungsten (ppm) | 0.52 | 0.43618 | 0.60382 | 0.44016 | 0.59984
Y, Yttrium (ppm) | 8.9951 | 8.4294 | 9.5608 | 8.7517 | 9.2384
Yb, Ytterbium (ppm) | 0.86718 | 0.6957 | 1.0386 | 0.78023 | 0.95412
Zn, Zinc (ppm) | 86.8375 | 84.0296 | 89.6454 | 85.1854 | 88.4896
Zr, Zirconium (ppm) | 21.2593 | 19.5585 | 22.96 | 20.5345 | 21.984
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 238b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 0.88682 | Pt | ppb | 0.58129
4-Acid Digestion
Au | ppm | 2.9079 | Ir | ppm | 0.00065 | Ru | ppm | 0.0033167
B | ppm | 28.6733 | Pd | ppb | 5.0083 | Se | ppm | 0.8
Ge | ppm | 0.19974 | Pt | ppb | 1.0917 | Te | ppm | 0.056329
Hg | ppm | 0.029867 | Rh | ppm | 0.00081667
Aqua Regia Digestion
B | ppm | 15.509 | Ir | ppm | < 0.0015 | Se | ppm | 0.29015
Dy | ppm | 2.0462 | Nd | ppm | 23.1412 | Si | wt.% | 0.39457
Er | ppm | 0.90444 | Pd | ppb | 30.05 | Sm | ppm | 4.275
Eu | ppm | 0.65371 | Pr | ppm | 6.2926 | Tb | ppm | 0.45794
Gd | ppm | 4.0345 | Pt | ppb | 2.3667 | Te | ppm | 0.020467
Hg | ppm | 0.015272 | Rh | ppm | 0.00066667 | Tm | ppm | 0.11163
Ho | ppm | 0.33768 | Ru | ppm | < 0.0009
Borate Fusion XRF
Al2O3 | wt.% | 14.645 | MgO | wt.% | 2.925 | SiO2 | wt.% | 67.385
CaO | wt.% | 1.195 | MnO | wt.% | 0.05 | SO3 | wt.% | 0.4355
Fe2O3 | wt.% | 5.88 | Na2O | wt.% | 0.945 | TiO2 | wt.% | 0.82
K2O | wt.% | 3.34 | P2O5 | wt.% | 0.1505
Thermogravimetry
LOI1000 | wt.% | 2.5
Infrared Combustion
C | wt.% | 0.105 | S | wt.% | 0.16667
Laser Ablation ICP-MS
Ag | ppm | 0.35 | Hf | ppm | 6.635 | Sm | ppm | 7.345
As | ppm | 697.5 | Ho | ppm | 1.175 | Sn | ppm | 3.8
Ba | ppm | 696.5 | In | ppm | < 0.05 | Sr | ppm | 117.5
Be | ppm | 3 | La | ppm | 41.15 | Ta | ppm | 1.215
Bi | ppm | 0.51 | Lu | ppm | 0.47 | Tb | ppm | 0.965
Cd | ppm | 0.2 | Mn | wt.% | 0.0421 | Te | ppm | < 0.2
Ce | ppm | 80.25 | Mo | ppm | 1 | Th | ppm | 15.4
Co | ppm | 18 | Nb | ppm | 15.95 | Ti | wt.% | 0.501
Cr | ppm | 150 | Nd | ppm | 36.95 | Tl | ppm | 0.4
Cs | ppm | 9.555 | Ni | ppm | 69 | Tm | ppm | 0.505
Cu | ppm | 31 | Pb | ppm | 26.5 | U | ppm | 3.12
Dy | ppm | 5.865 | Pr | ppm | 9.79 | V | ppm | 113
Er | ppm | 3.365 | Rb | ppm | 162 | W | ppm | 4
Eu | ppm | 1.355 | Re | ppm | < 0.01 | Y | ppm | 31
Ga | ppm | 19.75 | Sb | ppm | 663.5 | Yb | ppm | 3.215
Gd | ppm | 6.22 | Sc | ppm | 14.75 | Zn | ppm | 100
Ge | ppm | 1.55 | Se | ppm | < 5 | Zr | ppm | 233.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 238b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Raw*PA | Raw ~350g sample packed into 300cc jar analysed by PhotonAssay (high-power X-ray activation)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 238b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Hermosillo, Sonora, Mexico
Carsurin, Cikarang, West Java, Indonesia
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MinAnalytical Services, Perth, WA, Australia
MinAnalytical Services, Perth, WA, Australia
MSALABS, Vancouver, BC, Canada
On Site Laboratory Services, Bendigo, VIC, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Australia Mineral Services, Perth, WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Upscaled Metrics
ORE - Lab-Upscaled RSD Results for CRM: OREAS 238b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
2 | 1 | ANSLu | 1 | 1 | 1 | 219534 | 0.083092 | 3.2228 | 3.3296
2 | 1 | ANSLu | 1 | 20 | 20 | 219535 | 0.085616 | 3.393 | 3.3387
2 | 1 | ANSLu | 1 | 10 | 10 | 219536 | 0.083957 | 3.4611 | 3.3423
2 | 1 | ANSLu | 1 | 6 | 6 | 219537 | 0.083477 | 3.2785 | 3.3326
2 | 1 | ANSLu | 1 | 7 | 7 | 219538 | 0.084867 | 3.2555 | 3.3313
2 | 1 | ANSLu | 1 | 19 | 19 | 219539 | 0.085074 | 3.312 | 3.3344
2 | 1 | ANSLu | 1 | 9 | 9 | 219540 | 0.082761 | 3.4637 | 3.3424
2 | 1 | ANSLu | 1 | 2 | 2 | 219541 | 0.083256 | 3.2822 | 3.3328
2 | 1 | ANSLu | 1 | 5 | 5 | 219542 | 0.08421 | 3.292 | 3.3333
2 | 1 | ANSLu | 1 | 13 | 13 | 219543 | 0.082857 | 3.2733 | 3.3323
2 | 1 | ANSLu | 1 | 8 | 8 | 219544 | 0.086885 | 3.3095 | 3.3342
2 | 1 | ANSLu | 1 | 4 | 4 | 219545 | 0.086046 | 3.2153 | 3.3292
2 | 1 | ANSLu | 1 | 16 | 16 | 219546 | 0.086225 | 3.3789 | 3.3379
2 | 1 | ANSLu | 1 | 18 | 18 | 219547 | 0.085308 | 3.5519 | 3.3471
2 | 1 | ANSLu | 1 | 14 | 14 | 219548 | 0.085476 | 3.3975 | 3.3389
2 | 1 | ANSLu | 1 | 11 | 11 | 219549 | 0.088033 | 3.2849 | 3.3329
2 | 1 | ANSLu | 1 | 12 | 12 | 219550 | 0.086717 | 3.2696 | 3.3321
2 | 1 | ANSLu | 1 | 17 | 17 | 219551 | 0.08831 | 3.3617 | 3.337
2 | 1 | ANSLu | 1 | 3 | 3 | 219552 | 0.084483 | 3.3917 | 3.3386
2 | 1 | ANSLu | 1 | 15 | 15 | 219553 | 0.088449 | 3.3172 | 3.3346
Mean | [6 empty] | 0.085255 | 3.3356 | 3.3356
Median | [7 empty] | 3.3108 | 3.3343
SD | [7 empty] | 0.087599 | 0.0046698
RSD | [7 empty] | 2.626% | 0.14%
RSD Upscale Factor | [7 empty] | 0.053309
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 238b (Certified Value 3.08 ppm) | [32 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*MS | 2
|  |  | 0.085g | 25g | 50g | 50g | 40g | 30g | 40g | 30g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 15g | 50g | 30g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 25g | 25g | 3
| 1 | 1 | 3.2228 | 3.274 | 3.13 | 3.043 | 3.01 | 3.02 | 3.17 | 3.02 | 3.03 | 3.05 | 2.996 | 3.11 | 2.96 | 3.15 | 3.1 | 3.151 | 3.052 | 3.153 | 3.21 | 3.01 | 3.29 | 3 | 3.06 | 3.23 | 3.038 | 3.09 | 3.12 | 3.131 | 3.0912 | 3.17 | 3.1899 | 1
| 1 | 2 | 3.2822 | 3.114 | 3.05 | 3.04 | 2.95 | 2.99 | 3.15 | 3.2 | 3.02 | 3.05 | 3.143 | 3.14 | 2.94 | 3.18 | 3.11 | 3.187 | 3.0344 | 3.153 | 3.34 | 3.02 | 3.28 | 3.05 | 3.086 | 3.26 | 3.015 | 2.89 | 3.15 | 3.098 | 3.0624 | 3.11 | 2.98 | 3
| 1 | 3 | 3.3917 | 3.143 | 3.15 | 3.012 | 3.05 | 3.09 | 3.15 | 3.18 | 3.02 | 3.08 | 3.039 | 3.08 | 3.01 | 3.13 | 3.08 | 3.181 | 3.0184 | 3.043 | 3.37 | 3.06 | 3.22 | 2.94 | 3.066 | 3.27 | 3.009 | 3.02 | 3.115 | 3.166 | 3.0912 | 3.09 | 2.9662 | 16
| 1 | 4 | 3.2153 | 3.232 | 3.27 | 3.002 | 2.98 | 3 | 3.17 | 3.12 | 3.02 | 3.12 | 3.018 | 3.11 | 2.99 | 3.05 | 3.14 | 3.144 | 2.9552 | 2.979 | 3.38 | 3.092 | 3.24 | 3.03 | 3.036 | 3.18 | 3.027 | 2.93 | 3 | 3.123 | 3.0336 | 3.14 | 2.9599 | 3.0816
| 1 | 5 | 3.292 | 3.272 | 3.12 | 3.01 | 3.01 | 2.98 | 3.18 | 3.1 | 3.02 | 3.08 | 2.953 | 3.19 | 3.08 | 3.12 | 3.06 | 3.124 | 3.0326 | 2.998 | 3.31 | 3.04 | 3.21 | 3.05 | 3.078 | 3.22 | 3.051 | 3.03 | 3.06 | 3.076 | 3.0432 | 3.19 | 2.976 | 7
| 1 | 6 | 3.2785 | 3.203 | 3.1 | 3.084 | 3.04 | 2.96 | 3.14 | 3.04 | 3.02 | 3.12 | 2.968 | 3.12 | 2.84 | 3.17 | 3.12 | 3.148 | 2.9658 | 3.138 | 3.33 | 3.084 | 3.3 | 3.05 | 3.023 | 3.26 | 3.053 | 3.1 | 3.099 | 3.054 | 3.1104 | 3.11 | 2.9533
|  | 7 | 2.00: 3.2555
|  | 8 | 2.00: 3.3095
|  | 9 | 2.00: 3.4637
|  | 10 | 2.00: 3.4611
|  | 11 | 2.00: 3.2849
|  | 12 | 2.00: 3.2696
|  | 13 | 2.00: 3.2733
|  | 14 | 2.00: 3.3975
|  | 15 | 2.00: 3.3172
|  | 16 | 2.00: 3.3789
|  | 17 | 2.00: 3.3617
|  | 18 | 2.00: 3.5519
|  | 19 | 2.00: 3.312
|  | 20 | 2.00: 3.393
| Mean |  | 3.3356 | 3.2063 | 3.1367 | 3.0318 | 3.0067 | 3.0067 | 3.16 | 3.11 | 3.0217 | 3.0833 | 3.0195 | 3.125 | 2.97 | 3.1333 | 3.1017 | 3.1558 | 3.0097 | 3.0773 | 3.3233 | 3.051 | 3.2567 | 3.02 | 3.0582 | 3.2367 | 3.0322 | 3.01 | 3.0907 | 3.108 | 3.072 | 3.135 | 3.0042
| Median |  | 3.3108 | 3.2175 | 3.125 | 3.026 | 3.01 | 2.995 | 3.16 | 3.11 | 3.02 | 3.08 | 3.007 | 3.115 | 2.975 | 3.14 | 3.105 | 3.1495 | 3.0255 | 3.0905 | 3.335 | 3.05 | 3.26 | 3.04 | 3.063 | 3.245 | 3.0325 | 3.025 | 3.107 | 3.1105 | 3.0768 | 3.125 | 2.9711
| Std Dev. |  | 0.087599 | 0.066479 | 0.073666 | 0.030571 | 0.037238 | 0.045461 | 0.015492 | 0.072388 | 0.0040825 | 0.031411 | 0.068208 | 0.037283 | 0.08 | 0.046762 | 0.028577 | 0.023845 | 0.039755 | 0.080341 | 0.06121 | 0.033508 | 0.038297 | 0.043818 | 0.024334 | 0.033862 | 0.018335 | 0.084617 | 0.053298 | 0.040393 | 0.030358 | 0.038859 | 0.091507
| Rel.Std.Dev. |  | 2.626% | 2.073% | 2.349% | 1.008% | 1.239% | 1.512% | 0.4903% | 2.328% | 0.1351% | 1.019% | 2.259% | 1.193% | 2.694% | 1.492% | 0.9214% | 0.7556% | 1.321% | 2.611% | 1.842% | 1.098% | 1.176% | 1.451% | 0.7957% | 1.046% | 0.6047% | 2.811% | 1.724% | 1.3% | 0.9882% | 1.24% | 3.046%
| PDM3 |  | 8.244% | 4.049% | 1.788% | -1.614% | -2.431% | -2.431% | 2.545% | 0.9225% | -1.944% | 0.05713% | -2.014% | 1.409% | -3.621% | 1.68% | 0.6521% | 2.41% | -2.332% | -0.1376% | 7.845% | -0.9921% | 5.682% | -1.998% | -0.7596% | 5.033% | -1.603% | -2.323% | 0.2951% | 0.8576% | -0.3106% | 1.734% | -2.511%
| Z-Score (Absolute) |  | NA | 1.48 | 0.66 | 0.57 | 0.86 | 0.86 | 0.93 | 0.35 | 0.69 | 0.04 | 0.71 | 0.52 | 1.29 | 0.62 | 0.25 | 0.89 | 0.83 | 0.04 | 2.85 | 0.34 | 2.07 | 0.71 | 0.26 | 1.83 | 0.56 | 0.82 | 0.12 | 0.32 | 0.1 | 0.64 | 0.89
| Analytical results for Pd in OREAS 238b (Indicative Value 0.887 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.34 | ppb
|  |  | 2.00: FA*MS | 3
|  |  | 2.00: 25g | 3
| 1 | 1 | 2.00: 0.93433 | 1
| 1 | 2 | 2.00: 0.8855 | 1
| 1 | 3 | 2.00: 0.89033 | 16
| 1 | 4 | 2.00: 0.90948 | 0.88682
| 1 | 5 | 2.00: 0.83368 | 7
| 1 | 6 | 2.00: 0.86759
| Mean |  | 2.00: 0.88682
| Median |  | 2.00: 0.88792
| Std Dev. |  | 2.00: 0.034579
| Rel.Std.Dev. |  | 2.00: 3.899%
| PDM3 |  | 2.00: 2.22e-14%
Z-Score (Absolute) | bar: 2.00
| Analytical results for Pt in OREAS 238b (Indicative Value 0.581 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | 2.00: Lab | 1
| No | No | 2.00: 1.34 | ppb
|  |  | 2.00: FA*MS | 3
|  |  | 2.00: 25g | 3
| 1 | 1 | 2.00: 0.65489 | 1
| 1 | 2 | 2.00: 0.62819 | 1
| 1 | 3 | 2.00: 0.54911 | 16
| 1 | 4 | 2.00: 0.56789 | 0.58129
| 1 | 5 | 2.00: 0.5314 | 7
| 1 | 6 | 2.00: 0.55626
| Mean |  | 2.00: 0.58129
| Median |  | 2.00: 0.56207
| Std Dev. |  | 2.00: 0.048882
| Rel.Std.Dev. |  | 2.00: 8.409%
PDM3 | bar: 2.00
Z-Score (Absolute) | bar: 2.00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 238b (Certified Value 3 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.20 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | ppm
|  |  | INAA | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | 2
|  |  | 0.085g | 25g | 50g | 30g | 40g | 30g | 30g | 25g | 30g | 25g | 25g | 25g | 25g | 15g | 50g | 25g | 50g | 25g | 10g | 25g | 20g | 15g | 15g | 50g | 3
| 1 | 1 | 3.2228 | 2.95 | 2.88 | 3.0958 | 2.98 | > 2 | 3.02 | > 0.5 | 2.95 | 2.85 | 2.91 | 2.83 | 3.03 | 3.0212 | 2.84 | 3.19 | 2.99 | 2.56 | 3.04 | 3.062 | 3 | 2.7586 | 2.7938 | 3.0268 | 1
| 1 | 2 | 3.2822 | 3.02 | 3 | 3.0667 | 2.89 | > 2 | 3.03 | > 0.5 | 2.99 | 2.89 | 3.04 | 2.94 | 3.01 | 2.972 | 3.17 | 3.15 | 2.92 | 2.64 | 3.05 | 3.09 | 2.992 | 2.7896 | 2.7879 | 3.03 | 3
| 1 | 3 | 3.3917 | 3.13 | 3 | 3.1346 | 2.88 | > 2 | 3.02 | > 0.5 | 3 | 2.88 | 2.96 | 3.01 | 3.06 | 3.1094 | 3 | 3.17 | 2.97 | 2.72 | 3.09 | 3.059 | 2.9936 | 2.8308 | 2.8053 | 3.0578 | 16
| 1 | 4 | 3.2153 | 3.06 | 3.03 | 3.0376 | 2.9 | > 2 | 3.02 | > 0.5 | 2.96 | 2.9 | 3.04 | 2.85 | 3 | 3.0611 | 2.95 | 3.1 | 2.98 | 2.8 | 3.08 | 2.984 | 2.9899 | 2.772 | 2.9143 | 3.06 | 2.9986
| 1 | 5 | 3.292 | 3.01 | 2.94 | 2.9891 | 3.02 | > 2 | 3.04 | > 0.5 | 2.94 | 2.93 | 3.12 | 2.85 | 3.01 | 3.1698 | 3.09 | 3.12 | 2.94 | 2.42 | 3.07 | 2.99 | 2.9863 | 2.8447 | 2.9241 | 3.06 | 9
| 1 | 6 | 3.2785 | 3.05 | 3.02 | 3.0249 | 2.92 | > 2 | 3.05 | > 0.5 | 2.92 | 2.9 | 3.16 | 2.83 | 3.02 | 2.9638 | 2.83 | 3.12 | 2.97 | 2.6 | 3.08 | 3.11 | 2.967 | 2.7636 | 2.9534 | 3.04
|  | 7 | 3.2555
|  | 8 | 3.3095
|  | 9 | 3.4637
|  | 10 | 3.4611
|  | 11 | 3.2849
|  | 12 | 3.2696
|  | 13 | 3.2733
|  | 14 | 3.3975
|  | 15 | 3.3172
|  | 16 | 3.3789
|  | 17 | 3.3617
|  | 18 | 3.5519
|  | 19 | 3.312
|  | 20 | 3.393
| Mean |  | 3.3356 | 3.0367 | 2.9783 | 3.0581 | 2.9317 |  | 3.03 |  | 2.96 | 2.8917 | 3.0383 | 2.885 | 3.0217 | 3.0495 | 2.98 | 3.1417 | 2.9617 | 2.6233 | 3.0683 | 3.0492 | 2.9881 | 2.7932 | 2.8631 | 3.0458
| Median |  | 3.3108 | 3.035 | 3 | 3.0522 | 2.91 |  | 3.025 |  | 2.955 | 2.895 | 3.04 | 2.85 | 3.015 | 3.0412 | 2.975 | 3.135 | 2.97 | 2.62 | 3.075 | 3.0605 | 2.991 | 2.7808 | 2.8598 | 3.0489
| Std Dev. |  | 0.087599 | 0.059889 | 0.057417 | 0.052213 | 0.056006 |  | 0.012649 |  | 0.030332 | 0.026394 | 0.093897 | 0.073689 | 0.02137 | 0.080432 | 0.13535 | 0.034303 | 0.026394 | 0.13171 | 0.019408 | 0.051724 | 0.011303 | 0.036337 | 0.075226 | 0.015437
| Rel.Std.Dev. |  | 2.626% | 1.972% | 1.928% | 1.707% | 1.91% |  | 0.4175% |  | 1.025% | 0.9128% | 3.09% | 2.554% | 0.7072% | 2.637% | 4.542% | 1.092% | 0.8912% | 5.021% | 0.6325% | 1.696% | 0.3783% | 1.301% | 2.627% | 0.5068%
| PDM3 |  | 11.24% | 1.269% | -0.6761% | 1.985% | -2.232% |  | 1.047% |  | -1.287% | -3.566% | 1.325% | -3.789% | 0.769% | 1.699% | -0.6205% | 4.771% | -1.232% | -12.51% | 2.325% | 1.686% | -0.3493% | -6.85% | -4.518% | 1.573%
| Z-Score (Absolute) |  | NA | 0.57 | 0.11 | 0.82 | 0.66 | NA | 0.49 | NA | 0.33 | 1.13 | 0.59 | 1.21 | 0.39 | 0.72 | 0.1 | 1.8 | 0.31 | 4.27 | 0.94 | 0.71 | 0 | 2.28 | 1.46 | 0.67
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 238b (Certified Value 2.74 ppm) | [19 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 1.01 | 1.02 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.28 | 1.30 | ppm
|  |  | INAA | CNL*MS | CNL*AAS | CNL*AAS | CNL*MS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*AAS | CNL*AAS | CNL*OES | 2
|  |  | 0.085g | 200g | 200g | 200g | 50g | 50g | 20g | 200g | 05g | 30g | 30g | 30g | 30g | 30g | 200g | 50g | 30g | 15g | 3
| 1 | 1 | 3.2228 | 2.72 | 2.79 | 2.98 | 2.47 | 2.73 | 2.54 | 2.85 | 3 | 2.65 | 2.78 | 2.67 | 2.69 | 2.53 | 2.6022 | 3.01 | 2.2 | 2.793 | 1
| 1 | 2 | 3.2822 | 2.67 | 2.74 | 2.92 | 2.49 | 2.71 | 2.64 | 2.83 | 2.9 | 2.64 | 2.86 | 2.69 | 2.71 | 2.52 | 2.5535 | 2.967 | 2.22 | 2.8405 | 3
| 1 | 3 | 3.3917 | 2.72 | 2.84 | 2.93 | 2.46 | 2.75 | 2.6 | 2.84 | 3 | 2.54 | 3.14 | 2.71 | 2.71 | 2.56 | 2.5327 | 3.03 | 2.24 | 2.831 | 16
| 1 | 4 | 3.2153 | 2.72 | 2.78 | 3.04 | 2.48 | 2.75 | 2.55 | 2.84 | 3 | 2.6 | 3.02 | 2.71 | 2.76 | 2.52 | 2.5312 | 2.87 | 2.24 | 2.7835 | 2.7425
| 1 | 5 | 3.292 | 2.72 | 2.8 | 2.98 | 2.45 | 2.71 | 2.63 | 2.84 | 2.9 | 2.58 | 3.07 | 2.78 | 2.69 | 2.51 | 2.4825 | 2.87 | 2.21 | 2.85 | 11
| 1 | 6 | 3.2785 | 2.65 | 2.82 | 2.95 | 2.48 | 2.71 | 2.59 | 2.83 | 2.9 | 2.54 | 3.03 | 2.72 | 2.73 | 2.55 | 2.5917 | 2.87 | 2.21 | 2.8215
|  | 7 | 2.00: 3.2555
|  | 8 | 2.00: 3.3095
|  | 9 | 2.00: 3.4637
|  | 10 | 2.00: 3.4611
|  | 11 | 2.00: 3.2849
|  | 12 | 2.00: 3.2696
|  | 13 | 2.00: 3.2733
|  | 14 | 2.00: 3.3975
|  | 15 | 2.00: 3.3172
|  | 16 | 2.00: 3.3789
|  | 17 | 2.00: 3.3617
|  | 18 | 2.00: 3.5519
|  | 19 | 2.00: 3.312
|  | 20 | 2.00: 3.393
| Mean |  | 3.3356 | 2.7 | 2.795 | 2.9667 | 2.4717 | 2.7267 | 2.5917 | 2.8383 | 2.95 | 2.5917 | 2.9833 | 2.7133 | 2.715 | 2.5317 | 2.549 | 2.9362 | 2.22 | 2.8199
| Median |  | 3.3108 | 2.72 | 2.795 | 2.965 | 2.475 | 2.72 | 2.595 | 2.84 | 2.95 | 2.59 | 3.025 | 2.71 | 2.71 | 2.525 | 2.5431 | 2.9185 | 2.215 | 2.8262
| Std Dev. |  | 0.087599 | 0.031623 | 0.034496 | 0.043665 | 0.01472 | 0.019664 | 0.040702 | 0.0075277 | 0.054772 | 0.047504 | 0.13574 | 0.037238 | 0.026646 | 0.019408 | 0.04402 | 0.075287 | 0.016733 | 0.026475
| Rel.Std.Dev. |  | 2.626% | 1.171% | 1.234% | 1.472% | 0.5955% | 0.7212% | 1.571% | 0.2652% | 1.857% | 1.833% | 4.55% | 1.372% | 0.9814% | 0.7666% | 1.727% | 2.564% | 0.7537% | 0.9389%
| PDM3 |  | 21.63% | -1.55% | 1.914% | 8.174% | -9.875% | -0.5774% | -5.5% | 3.494% | 7.566% | -5.5% | 8.781% | -1.064% | -1.003% | -7.688% | -7.057% | 7.062% | -19.05% | 2.823%
| Z-Score (Absolute) |  | NA | 0.08 | 0.44 | 1.38 | 1.33 | 0.06 | 0.67 | 0.67 | 1.28 | 0.67 | 1.47 | 0.01 | 0 | 1 | 0.91 | 1.21 | 2.71 | 0.57
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PA
| Analytical results for Au in OREAS 238b (Certified Value 3.16 ppm) | [6 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 2.00 | 5.01 | 5.02 | 5.03 | 5.04 | ppm
|  |  | INAA | Raw*PA | Raw*PA | Raw*PA | Raw*PA | 2
|  |  | 0.085g | 350g | 350g | 350g | 350g | 3
| 1 | 1 | 3.2228 | 3.11 | 3.18 | 3.15 | 3.02 | 1
| 1 | 2 | 3.2822 | 3.07 | 3.19 | 3.15 | 3.34 | 3
| 1 | 3 | 3.3917 | 3.13 | 3.13 | 3.15 | 3.26 | 16
| 1 | 4 | 3.2153 | 3.17 | 3.12 | 3.15 | 3.12 | 3.164
| 1 | 5 | 3.292 | 2.9 | 3.18 | 3.15 | 3.14 | 13
| 1 | 6 | 3.2785 | 3.04 | 3.04 | 3.15 | 3.14
| 2 | 7 | 3.2555 | 3.181 | 3.07 | 3.31 | 3.25
| 2 | 8 | 3.3095 | 3.119 | 3.11 | 3.31 | 3.27
| 2 | 9 | 3.4637 | 2.906 | 3.23 | 3.31 | 3.22
| 2 | 10 | 3.4611 | 3.254 | 2.98 | 3.31 | 3.12
| 2 | 11 | 3.2849 | 3.627 | 2.93 | 3.31 | 3.26
| 2 | 12 | 3.2696 | 3.173 | 3.15 | 3.31 | 3.17
| 3 | 13 | 3.2733 |  | 3.13 | 3.19
| 3 | 14 | 3.3975 |  | 3.15 | 3.19
| 3 | 15 | 3.3172 |  | 3.16 | 3.19
| 3 | 16 | 3.3789 |  | 3.14 | 3.19
| 3 | 17 | 3.3617 |  | 3.22 | 3.19
| 3 | 18 | 3.5519 |  | 3.15 | 3.19
| 4 | 19 | 3.312 |  | 3.084 | 3.166
| 4 | 20 | 3.393 |  | 3.185 | 3.101
| 4 | 21 |  |  | 3.163 | 3.18
| 4 | 22 |  |  | 3.061 | 3.255
| 4 | 23 |  |  | 3.184 | 3.18
| 4 | 24 |  |  | 3.153 | 3.127
| 5 | 25 |  |  | 3.138 | 3.011
| 5 | 26 |  |  | 3.196 | 3.044
| 5 | 27 |  |  | 3.261 | 3.199
| 5 | 28 |  |  | 3.215 | 3.154
| 5 | 29 |  |  | 3.199 | 3.352
| 5 | 30 |  |  | 3.085 | 3.152
| 6 | 31 |  |  | 3.1459 | 3.215
| 6 | 32 |  |  | 3.3174 | 3.37
| 6 | 33 |  |  | 3.1296 | 3.395
| 6 | 34 |  |  | 3.2217 | 3.263
| 6 | 35 |  |  | 3.3701 | 3.254
| 6 | 36 |  |  | 3.2928 | 3.258
| Mean |  | 3.3356 | 3.14 | 3.1573 | 3.2104 | 3.1925
| Median |  | 3.3108 | 3.1245 | 3.1515 | 3.19 | 3.195
| Std Dev. |  | 0.087599 | 0.18648 | 0.085322 | 0.087728 | 0.089048
| Rel.Std.Dev. |  | 2.626% | 5.939% | 2.702% | 2.733% | 2.789%
| PDM3 |  | 5.425% | -0.758% | -0.2124% | 1.468% | 0.9013%
| Z-Score (Absolute) |  | NA | 0.9 | 0.45 | 0.91 | 0.45
Note on C47: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 238b (Certified Value 0.245 ppm) | [27 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.26 | 0.2 | 0.24 | 0.7282 | 0.3 | < 0.5 | 0.4 | 0.5 | 0.23 | 0.23 | 0.23 | 0.25 | 0.18 | < 1 | 0.41 | 0.37 | < 0.5 | 0.28 | 0.25 | 0.206 | 0.36702 | 1.0671 |  |  |  |  | 1
| 1 | 2 | 0.26 | 0.2 | 0.25 | 0.7194 | 0.2 | < 0.5 | 0.3 | 0.5 | 0.23 | 0.23 | 0.25 | 0.25 | 0.21 | < 1 | 0.41 | 0.34 | < 0.5 | 0.27 | 0.23 | 0.277 | 0.37773 | 1.3175 |  |  |  |  | 25
| 1 | 3 | 0.26 | 0.2 | 0.27 | 0.721 | 0.2 | < 0.5 | 0.3 | 0.4 | 0.23 | 0.22 | 0.24 | 0.25 | 0.13 | < 1 | 0.41 | 0.32 | < 0.5 | 0.26 | 0.24 | 0.207 | 0.41601 | 1.4024 |  |  |  |  | 16
| 1 | 4 | 0.26 | 0.2 | 0.25 | 0.71834 | 0.3 | < 0.5 | 0.4 | 0.5 | 0.24 | 0.25 | 0.26 | 0.25 | 0.18 | < 1 | 0.41 | 0.33 | < 0.5 | 0.2 | 0.21 | 0.219 | 0.33761 | 1.2793 |  |  |  |  | 0.24462
| 1 | 5 | 0.26 | 0.2 | 0.29 | 0.7076 | 0.3 | < 0.5 | 0.4 | 0.3 | 0.23 | 0.24 | 0.25 | 0.24 | 0.17 | < 1 | 0.41 | 0.32 | < 0.5 | 0.24 | 0.24 | 0.213 | 0.33489 | 1.3175 |  |  |  |  | 15
| 1 | 6 | 0.26 | 0.2 | 0.25 | 0.73 | 0.2 | < 0.5 | 0.3 | 0.3 | 0.24 | 0.25 | 0.25 | 0.27 | 0.2 | < 1 | 0.41 | 0.33 | < 0.5 | 0.23 | 0.25 | 0.211 | 0.34717 | 1.169
| Mean |  | 0.26 | 0.2 | 0.25833 | 0.72076 | 0.25 |  | 0.35 | 0.41667 | 0.23333 | 0.23667 | 0.24667 | 0.25167 | 0.17833 |  | 0.41 | 0.335 |  | 0.24667 | 0.23667 | 0.22217 | 0.3634 | 1.2588
| Median |  | 0.26 | 0.2 | 0.25 | 0.7202 | 0.25 |  | 0.35 | 0.45 | 0.23 | 0.235 | 0.25 | 0.25 | 0.18 |  | 0.41 | 0.33 |  | 0.25 | 0.24 | 0.212 | 0.35709 | 1.2984
| Std Dev. |  | 0 | 3.0405e-17 | 0.018348 | 0.0080212 | 0.054772 |  | 0.054772 | 0.098319 | 0.005164 | 0.012111 | 0.010328 | 0.0098319 | 0.027869 |  | 0 | 0.018708 |  | 0.029439 | 0.015055 | 0.027265 | 0.030749 | 0.12062
| Rel.Std.Dev. |  | 0% | 1.52e-14% | 7.103% | 1.113% | 21.91% |  | 15.65% | 23.6% | 2.213% | 5.117% | 4.187% | 3.907% | 15.63% |  | 0% | 5.585% |  | 11.93% | 6.361% | 12.27% | 8.461% | 9.582%
| PDM3 |  | 6.285% | -18.24% | 5.604% | 194.6% | 2.198% |  | 43.08% | 70.33% | -4.616% | -3.253% | 0.8349% | 2.879% | -27.1% |  | 67.6% | 36.94% |  | 0.8349% | -3.253% | -9.18% | 48.56% | 414.6%
| Z-Score (Absolute) |  | 0.39 | NA | 0.32 | 20.01 | NA | 0.04 | NA | NA | 0.75 | 0.6 | 0.18 | 0.04 | 3.09 | 10.61 | 6.78 | 3.58 | 0.04 | 0.18 | 0.6 | 1.22 | 4.79 | 42.93
Results from laboratories 1.05, 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 238b (Certified Value 7.47 wt.%) | [27 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 7.4694 | 7.4618 | 7.54 | 7.8 | 7.56 | 7.513 | 6.87 | 7.5 | 7.2 | 7.29 | 7.34 | 7.63 | 7.53 | 7.6275 | 7.1321 | 7.515 | 7.55 | 7.37 | 7.32 | 7.5 | 6.89 | 7.176 | 8.0147 | 7.4765 |  |  | 1
| 1 | 2 | 7.4261 | 7.5398 | 7.54 | 7.97 | 7.61 | 7.596 | 6.51 | 7.64 | 7.11 | 7.33 | 7.55 | 7.49 | 7.41 | 7.0629 | 7.1031 | 7.727 | 7.63 | 7.41 | 7.27 | 7.58 | 6.87 | 7.1055 | 8.0249 | 7.4812 |  |  | 0
| 1 | 3 | 7.4715 | 7.496 | 7.44 | 7.99 | 7.6336 | 7.625 | 7.15 | 7.54 | 7.27 | 7.38 | 7.57 | 7.56 | 7.28 | 6.9397 | 7.098 | 8.68 | 7.43 | 7.42 | 7.29 | 7.57 | 7.03 | 7.3531 | 7.9525 | 7.4654 |  |  | 16
| 1 | 4 | 7.6089 | 7.4883 | 7.44 | 8.03 | 7.56 | 7.493 | 8.05 | 7.71 | 7.33 | 7.61 | 7.38 | 7.55 | 7.47 | 7.2205 | 7.194 | 7.515 | 7.49 | 7.31 | 7.3 | 7.48 | 7.07 | 7.2094 | 8.3776 | 7.5017 |  |  | 7.4723
| 1 | 5 | 7.5736 | 7.4898 | 7.57 | 7.9 | 7.63 | 7.569 | 7.71 | 7.42 | 7.15 | 7.52 | 7.28 | 7.53 | 7.71 | 6.965 | 7.2361 | 7.304 | 7.5 | 7.35 | 6.91 | 7.49 | 6.91 | 7.2341 | 7.9568 | 7.4955 |  |  | 16
| 1 | 6 | 7.6223 | 7.5017 | 7.5 | 7.56 | 7.6032 | 7.599 | 7.18 | 7.6 | 7.29 | 7.8 | 7.4 | 7.71 | 7.61 | 6.8458 | 7.404 | 7.515 | 7.4 | 7.36 | 7.21 | 7.55 | 6.92 | 7.141 | 7.7926 | 7.4947
| Mean |  | 7.5286 | 7.4962 | 7.505 | 7.875 | 7.5995 | 7.5658 | 7.245 | 7.5683 | 7.225 | 7.4883 | 7.42 | 7.5783 | 7.5017 | 7.1102 | 7.1946 | 7.7093 | 7.5 | 7.37 | 7.2167 | 7.5283 | 6.9483 | 7.2032 | 8.0198 | 7.4858
| Median |  | 7.5225 | 7.4929 | 7.52 | 7.935 | 7.6066 | 7.5825 | 7.165 | 7.57 | 7.235 | 7.45 | 7.39 | 7.555 | 7.5 | 7.014 | 7.1631 | 7.515 | 7.495 | 7.365 | 7.28 | 7.525 | 6.915 | 7.1927 | 7.9858 | 7.4879
| Std Dev. |  | 0.083095 | 0.025368 | 0.055045 | 0.17421 | 0.032673 | 0.052186 | 0.55827 | 0.10362 | 0.085732 | 0.19447 | 0.1161 | 0.079099 | 0.15132 | 0.28357 | 0.1159 | 0.49398 | 0.082946 | 0.040497 | 0.15488 | 0.043551 | 0.081588 | 0.086738 | 0.19401 | 0.013783
| Rel.Std.Dev. |  | 1.104% | 0.3384% | 0.7335% | 2.212% | 0.4299% | 0.6898% | 7.706% | 1.369% | 1.187% | 2.597% | 1.565% | 1.044% | 2.017% | 3.988% | 1.611% | 6.408% | 1.106% | 0.5495% | 2.146% | 0.5785% | 1.174% | 1.204% | 2.419% | 0.1841%
| PDM3 |  | 0.7543% | 0.3206% | 0.438% | 5.39% | 1.702% | 1.252% | -3.042% | 1.286% | -3.309% | 0.215% | -0.6995% | 1.419% | 0.3934% | -4.845% | -3.717% | 3.173% | 0.3711% | -1.369% | -3.421% | 0.7503% | -7.012% | -3.601% | 7.328% | 0.1815%
| Z-Score (Absolute) |  | 0.23 | 0.01 | 0.05 | 2.8 | 0.75 | 0.5 | 1.88 | 0.52 | 2.03 | 0.07 | 0.58 | 0.6 | 0.03 | 2.88 | 2.25 | 1.57 | 0.01 | 0.95 | 2.09 | 0.22 | 4.08 | 2.19 | 3.88 | 0.09
| Analytical results for As in OREAS 238b (Certified Value 676 ppm) | [27 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 0
0
| 1 | 1 | 675.7 | 642 | 687 | 677.99 | 665 | 661 | 622 | 216.5 | 660 | 681 | 690 | 709 | 705.5 | 623.3098 | 747 | 652.7 | 661.7 | 703.7 | 276.6 | 650.1 | 825.9773 | [5 empty] | 1
| 1 | 2 | 680.1 | 641 | 730 | 676.49 | 661 | 682 | 623 | 207.1 | 670 | 694 | 633 | 697 | 690.9 | 624.3697 | 751 | 645 | 639.8 | 714.6 | 281.6 | 639.8 | 697.4973 | [5 empty] | 26
| 1 | 3 | 672.5 | 643 | 699 | 680.2898 | 676 | 665 | 667 | 205.9 | 689 | 707 | 651 | 689 | 697.4 | 634.8676 | 713 | 658.1 | 668.2 | 724.5 | 292.4 | 655.8 | 696.7659 | [5 empty] | 16
| 1 | 4 | 688.6 | 629 | 705 | 676.5556 | 674 | 691 | 667 | 214.7 | 719 | 690 | 655 | 700 | 704.7 | 615.8168 | 729 | 652.9 | 659.4 | 712 | 278.6 | 651.6 | 661.4766 | [5 empty] | 676.0572
| 1 | 5 | 676.4 | 641 | 699 | 680.14 | 665 | 674 | 606 | 208.9 | 706 | 682 | 661 | 719 | 693.5 | 635.5172 | 749 | 661.4 | 659.1 | 704 | 261.7 | 653.3 | 630.2912 | [5 empty] | 17
| 1 | 6 | 688.3 | 639 | 718 | 675.65 | 668 | 688 | 646 | 204.8 | 715 | 685 | 666 | 710 | 695.6 | 622.2159 | 710 | 652.1 | 637.5 | 706 | 318.8 | 637.8 | 752.3105
| Mean |  | 680.2667 | 639.1667 | 706.3333 | 677.8526 | 668.1667 | 676.8333 | 638.5 | 209.65 | 693.1667 | 689.8333 | 659.3333 | 704 | 697.9333 | 626.0162 | 733.1667 | 653.7 | 654.2833 | 710.8 | 284.95 | 648.0667 | 710.7198
| Median |  | 678.25 | 641 | 702 | 677.2728 | 666.5 | 678 | 634.5 | 208 | 697.5 | 687.5 | 658 | 704.5 | 696.5 | 623.8398 | 738 | 652.8 | 659.25 | 709 | 280.1 | 650.85 | 697.1316
| Std Dev. |  | 6.7843 | 5.1543 | 15.358 | 1.979 | 5.7764 | 12.2543 | 25.4931 | 4.8389 | 24.3427 | 9.7451 | 18.8113 | 10.7331 | 5.9648 | 7.7087 | 18.5517 | 5.631 | 12.5659 | 8.0406 | 19.3037 | 7.4505 | 69.6811
| Rel.Std.Dev. |  | 0.9973% | 0.8064% | 2.174% | 0.2919% | 0.8645% | 1.811% | 3.993% | 2.308% | 3.512% | 1.413% | 2.853% | 1.525% | 0.8546% | 1.231% | 2.53% | 0.8614% | 1.921% | 1.131% | 6.774% | 1.15% | 9.804%
| PDM3 |  | 0.6226% | -5.457% | 4.478% | 0.2656% | -1.167% | 0.1148% | -5.555% | -68.99% | 2.531% | 2.038% | -2.474% | 4.133% | 3.236% | -7.402% | 8.447% | -3.307% | -3.221% | 5.139% | -57.85% | -4.14% | 5.127%
| Z-Score (Absolute) |  | 0.09 | 0.93 | 0.73 | 0.03 | 0.22 | 0 | 0.95 | 11.6 | 0.41 | 0.32 | 0.43 | 0.67 | 0.52 | 1.26 | 1.4 | 0.57 | 0.56 | 0.84 | 9.73 | 0.71 | 0.84
| Analytical results for Au in OREAS 238b (Indicative Value 2.91 ppm) | [27 empty] | Indicative
Au | Round | Replicate | 1.01: Lab | [5 empty] | 1
| No | No | 1.01: 1.30 | [5 empty] | ppm
|  |  | 1.01: 4A*MS | [5 empty] | 2
2
| 1 | 1 | 1.01: 2.7346 | [5 empty] | 1
| 1 | 2 | 1.01: 3.0846 | [5 empty] | 3
| 1 | 3 | 1.01: 2.9957 | [5 empty] | 16
| 1 | 4 | 1.01: 2.8332 | [5 empty] | 2.9079
| 1 | 5 | 1.01: 2.9768 | [5 empty] | 9
| 1 | 6 | 1.01: 2.8226
| Mean |  | 1.01: 2.9079
| Median |  | 1.01: 2.905
| Std Dev. |  | 1.01: 0.13158
| Rel.Std.Dev. |  | 1.01: 4.525%
| PDM3 |  | 1.01: 2.22e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for B in OREAS 238b (Indicative Value 28.7 ppm) | [27 empty] | Indicative
B | Round | Replicate | 1.01: Lab | [5 empty] | 1
| No | No | 1.01: 1.07 | [5 empty] | ppm
|  |  | 1.01: 4A*OES | [5 empty] | 1
1
| 1 | 1 | 1.01: 30.84 | [5 empty] | 1
| 1 | 2 | 1.01: 27.08 | [5 empty] | 4
| 1 | 3 | 1.01: 30.35 | [5 empty] | 16
| 1 | 4 | 1.01: 29.45 | [5 empty] | 28.6733
| 1 | 5 | 1.01: 26.75 | [5 empty] | 10
| 1 | 6 | 1.01: 27.57
| Mean |  | 1.01: 28.6733
| Median |  | 1.01: 28.51
| Std Dev. |  | 1.01: 1.7643
| Rel.Std.Dev. |  | 1.01: 6.153%
| PDM3 |  | 1.01: 1.332e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ba in OREAS 238b (Certified Value 685 ppm) | [27 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 677.8 | 696.7218 | 705 | 748 | 672.2 | 717 | 686 | 717 | 636 | 690 | 680 | 700 | 700 | 704 | 675.4567 | 696 | 639.5 | 702 | 663.1 | 668 | 675 | 651 | 651.3 | 723 | 692.2615 | 608.1208 | 1
| 1 | 2 | 667.6 | 698.4248 | 699 | 743 | 675.14 | 727 | 697 | 715 | 654 | 700 | 700 | 680 | 680 | 669 | 676.0831 | 703 | 638.6 | 703 | 670.3 | 684 | 681 | 657 | 642.7 | 730 | 705.8117 | 585.4523 | 27
| 1 | 3 | 692 | 701.7007 | 694 | 768 | 669.6 | 726 | 696 | 712 | 659 | 700 | 700 | 690 | 670 | 681 | 685.5434 | 769 | 665.3 | 680 | 659.7 | 688 | 681 | 667 | 660.3 | 726 | 731.4393 | 591.0234 | 16
| 1 | 4 | 684.9 | 693.5124 | 699 | 762 | 674.77 | 711 | 711 | 725 | 646 | 710 | 680 | 690 | 690 | 686 | 673.6493 | 683 | 652.5 | 685 | 674.5 | 692 | 680 | 674 | 653.4 | 756 | 693.3907 | 582.8996 | 685.46
| 1 | 5 | 686.4 | 693.0742 | 708 | 742 | 676.2789 | 715 | 704 | 671 | 653 | 710 | 680 | 690 | 710 | 678 | 671.4592 | 669 | 682.4 | 689 | 668.3 | 673 | 677 | 651 | 656.4 | 719 | 695.9927 | 562.6269 | 18
| 1 | 6 | 693.8 | 692.5309 | 712 | 760 | 676.863 | 710 | 697 | 699 | 645 | 720 | 690 | 700 | 700 | 659 | 678.6126 | 691 | 645.7 | 671 | 667.6 | 671 | 681 | 645 | 644.8 | 704 | 696.9746 | 590.9518
| Mean |  | 683.75 | 695.9941 | 702.8333 | 753.8333 | 674.142 | 717.6667 | 698.5 | 706.5 | 648.8333 | 705 | 688.3333 | 691.6667 | 691.6667 | 679.5 | 676.8007 | 701.8333 | 654 | 688.3333 | 667.25 | 679.3333 | 679.1667 | 657.5 | 651.4833 | 726.3333 | 702.6451 | 586.8458
| Median |  | 685.65 | 695.1171 | 702 | 754 | 674.955 | 716 | 697 | 713.5 | 649.5 | 705 | 685 | 690 | 695 | 679.5 | 675.7699 | 693.5 | 649.1 | 687 | 667.95 | 678.5 | 680.5 | 654 | 652.35 | 724.5 | 696.4837 | 588.202
| Std Dev. |  | 9.7328 | 3.6244 | 6.6758 | 10.9255 | 2.7482 | 7.3121 | 8.4083 | 19.3468 | 8.1833 | 10.4881 | 9.8319 | 7.5277 | 14.7196 | 15.346 | 4.9092 | 34.9194 | 17.0458 | 12.5167 | 5.2405 | 9.9532 | 2.5626 | 10.9864 | 6.7443 | 17.0724 | 14.893 | 14.7716
| Rel.Std.Dev. |  | 1.423% | 0.5207% | 0.9498% | 1.449% | 0.4077% | 1.019% | 1.204% | 2.738% | 1.261% | 1.488% | 1.428% | 1.088% | 2.128% | 2.258% | 0.7254% | 4.975% | 2.606% | 1.818% | 0.7854% | 1.465% | 0.3773% | 1.671% | 1.035% | 2.35% | 2.12% | 2.517%
| PDM3 |  | -0.2495% | 1.537% | 2.535% | 9.975% | -1.651% | 4.699% | 1.902% | 3.069% | -5.343% | 2.851% | 0.4192% | 0.9055% | 0.9055% | -0.8695% | -1.263% | 2.389% | -4.59% | 0.4192% | -2.657% | -0.8938% | -0.9181% | -4.079% | -4.957% | 5.963% | 2.507% | -14.39%
| Z-Score (Absolute) |  | 0.22 | 0.36 | 0.69 | 3.1 | 0.67 | 1.39 | 0.48 | 0.86 | 1.87 | 0.79 | 0 | 0.16 | 0.16 | 0.42 | 0.55 | 0.64 | 1.62 | 0 | 1 | 0.43 | 0.43 | 1.46 | 1.74 | 1.8 | 0.68 | 4.8
| Analytical results for Be in OREAS 238b (Certified Value 2.47 ppm) | [27 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.53 | 2.5 | 2.6 | < 5 | 2.1 | 2.47 | 2.7 | 2 | 2.33 | 2.6 | 2.56 | 2.27 | 3 | 2.3838 | 3 | 2.042 | 2.8 | 3.17 | 2.45 | 2.35 | 2 | 2.3232 | 2.8311 | 2.4007 |  |  | 1
| 1 | 2 | 2.45 | 2.5 | 2.5 | < 5 | 2.1 | 2.57 | 2.6 | 2 | 2.39 | 2.53 | 2.47 | 2.27 | 2 | 2.3597 | 2.9 | 2.053 | 2.8 | 3.06 | 2.06 | 2.12 | 2 | 2.2924 | 2.998 | 2.431 |  |  | 28
| 1 | 3 | 2.42 | 2.5 | 2.5 | < 5 | 2.1 | 2.53 | 2.7 | 2 | 2.52 | 2.64 | 2.48 | 2.31 | 3 | 2.3568 | 2.9 | 2.112 | 2.5 | 3.15 | 2.35 | 2.08 | 2 | 2.3551 | 2.9424 | 2.4258 |  |  | 16
| 1 | 4 | 2.4 | 2.5 | 2.6 | < 5 | 2.4 | 2.67 | 2.5 | 2 | 2.57 | 2.58 | 2.61 | 2.32 | 2 | 2.2761 | 2.9 | 2.09 | 2.6 | 3.18 | 2.72 | 1.96 | 2 | 2.3331 | 2.7755 | 2.4261 |  |  | 2.4659
| 1 | 5 | 2.53 | 2.5 | 2.5 | < 5 | 2.2 | 2.63 | 2.2 | 2 | 2.41 | 2.54 | 2.39 | 2.44 | 3 | 2.3252 | 2.9 | 2.212 | 2.6 | 3.15 | 2.48 | 2.19 | 2 | 2.3419 | 2.8867 | 2.3531 |  |  | 19
| 1 | 6 | 2.49 | 2.5 | 2.6 | < 5 | 2.4 | 2.66 | 2.8 | 2 | 2.53 | 2.38 | 2.55 | 2.37 | 3 | 2.2603 | 2.8 | 2.088 | 2.5 | 3.1 | 2.38 | 2.15 | 2 | 2.2968 | 2.8867 | 2.3908
| Mean |  | 2.47 | 2.5 | 2.55 |  | 2.2167 | 2.5883 | 2.5833 | 2 | 2.4583 | 2.545 | 2.51 | 2.33 | 2.6667 | 2.327 | 2.9 | 2.0995 | 2.6333 | 3.135 | 2.4067 | 2.1417 | 2 | 2.3237 | 2.8867 | 2.4046
| Median |  | 2.47 | 2.5 | 2.55 |  | 2.15 | 2.6 | 2.65 | 2 | 2.465 | 2.56 | 2.515 | 2.315 | 3 | 2.341 | 2.9 | 2.089 | 2.6 | 3.15 | 2.415 | 2.135 | 2 | 2.3281 | 2.8867 | 2.4132
| Std Dev. |  | 0.055498 | 0 | 0.054772 |  | 0.1472 | 0.079099 | 0.2137 | 0 | 0.094745 | 0.090277 | 0.07874 | 0.065422 | 0.5164 | 0.04943 | 0.063246 | 0.060833 | 0.13663 | 0.045935 | 0.21426 | 0.12891 | 0 | 0.024936 | 0.078673 | 0.0299
| Rel.Std.Dev. |  | 2.247% | 0% | 2.148% |  | 6.64% | 3.056% | 8.272% | 0% | 3.854% | 3.547% | 3.137% | 2.808% | 19.36% | 2.124% | 2.181% | 2.898% | 5.188% | 1.465% | 8.903% | 6.019% | 0% | 1.073% | 2.725% | 1.243%
| PDM3 |  | 0.1657% | 1.382% | 3.41% |  | -10.11% | 4.964% | 4.762% | -18.89% | -0.3074% | 3.207% | 1.788% | -5.512% | 8.141% | -5.634% | 17.6% | -14.86% | 6.789% | 27.13% | -2.403% | -13.15% | -18.89% | -5.765% | 17.07% | -2.488%
| Z-Score (Absolute) |  | 0.21 | 0 | 0.35 | 0 | 2 | 0.62 | 0.59 | NA | 0.29 | 0.32 | 0.07 | 1.2 | NA | 1.22 | 2.83 | 2.83 | 0.94 | 4.49 | 0.66 | 2.53 | NA | 1.25 | 2.73 | 0.67
Results from laboratories 1.11, 1.16 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 238b (Certified Value 0.4 ppm) | [27 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 0.39 | 0.43653 | 0.4 | 0.37 | < 5 | 0.42 | 0.36 | 0.51 | 0.4 | 0.37 | 0.38 | 0.4 | 0.4 | 0.36 | 0.39934 | 0.5 | 0.412 | 0.35 | < 5 | 0.41 | 0.35 | 0.37 | 0.4608 | < 5 | 1.2955 | 0.38219 | 1
| 1 | 2 | 0.4 | 0.41634 | 0.4 | 0.39 | < 5 | 0.39 | 0.42 | 0.38 | 0.4 | 0.37 | 0.39 | 0.44 | 0.39 | 0.34 | 0.40747 | 0.5 | 0.358 | 0.41 | < 5 | 0.4 | 0.31 | 0.37 | 0.4936 | < 5 | 1.9062 | 0.37172 | 29
| 1 | 3 | 0.46 | 0.42041 | 0.4 | 0.39 | < 5 | 0.42 | 0.43 | 0.45 | 0.5 | 0.42 | 0.42 | 0.39 | 0.36 | 0.38 | 0.40972 | 0.5 | 0.373 | 0.4 | < 5 | 0.42 | 0.32 | 0.39 | 0.437 | < 5 | 1.3034 | 0.37864 | 16
| 1 | 4 | 0.4 | 0.4444 | 0.4 | 0.41 | < 5 | 0.38 | 0.44 | 0.44 | 0.5 | 0.42 | 0.39 | 0.41 | 0.37 | 0.35 | 0.39504 | 0.5 | 0.389 | 0.38 | < 5 | 0.39 | 0.3 | 0.39 | 0.4492 | < 5 | 1.2246 | 0.4064 | 0.40028
| 1 | 5 | 0.41 | 0.43221 | 0.4 | 0.38 | < 5 | 0.38 | 0.43 | 0.4 | 0.5 | 0.4 | 0.4 | 0.57 | 0.4 | 0.49 | 0.4031 | 0.5 | 0.404 | 0.38 | < 5 | 0.39 | 0.34 | 0.39 | 0.4271 | < 5 | 1.5891 | 0.41989 | 20
| 1 | 6 | 0.41 | 0.41917 | 0.4 | 0.38 | < 5 | 0.39 | 0.47 | 0.44 | 0.4 | 0.4 | 0.42 | 0.42 | 0.38 | 0.36 | 0.40851 | 0.5 | 0.381 | 0.37 | < 5 | 0.57 | 0.34 | 0.36 | 0.4143 | < 5 | 1.3862 | 0.39156
| Mean |  | 0.41167 | 0.42818 | 0.4 | 0.38667 |  | 0.39667 | 0.425 | 0.43667 | 0.45 | 0.39667 | 0.4 | 0.43833 | 0.38333 | 0.38 | 0.40386 | 0.5 | 0.38617 | 0.38167 |  | 0.43 | 0.32667 | 0.37833 | 0.447 |  | 1.4508 | 0.39173
| Median |  | 0.405 | 0.42631 | 0.4 | 0.385 |  | 0.39 | 0.43 | 0.44 | 0.45 | 0.4 | 0.395 | 0.415 | 0.385 | 0.36 | 0.40528 | 0.5 | 0.385 | 0.38 |  | 0.405 | 0.33 | 0.38 | 0.4431 |  | 1.3448 | 0.38688
| Std Dev. |  | 0.024833 | 0.011229 | 6.0809e-17 | 0.013663 |  | 0.018619 | 0.036194 | 0.045019 | 0.054772 | 0.022509 | 0.016733 | 0.066758 | 0.01633 | 0.055498 | 0.0057948 | 0 | 0.019934 | 0.02137 |  | 0.06957 | 0.019664 | 0.013292 | 0.028043 |  | 0.25601 | 0.018281
| Rel.Std.Dev. |  | 6.032% | 2.623% | 1.52e-14% | 3.533% |  | 4.694% | 8.516% | 10.31% | 12.17% | 5.675% | 4.183% | 15.23% | 4.26% | 14.6% | 1.435% | 0% | 5.162% | 5.599% |  | 16.18% | 6.02% | 3.513% | 6.274% |  | 17.65% | 4.667%
| PDM3 |  | 2.846% | 6.971% | -0.06875% | -3.4% |  | -0.9015% | 6.177% | 9.092% | 12.42% | -0.9015% | -0.06875% | 9.508% | -4.233% | -5.065% | 0.8965% | 24.91% | -3.525% | -4.649% |  | 7.426% | -18.39% | -5.482% | 11.67% |  | 262.5% | -2.134%
| Z-Score (Absolute) |  | 0.22 | 0.69 | NA | 0.49 | 59.23 | 0.2 | 0.6 | 0.93 | NA | 0.2 | 0.11 | 0.97 | 0.58 | 0.67 | 0 | NA | 0.5 | 0.63 | 59.23 | 0.74 | 2.18 | 0.72 | 1.22 | 59.23 | 29.58 | 0.34
Results from laboratories 1.05, 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 238b (Certified Value 0.855 wt.%) | [27 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.8531 | 0.84542 | 0.84 | 0.87 | 0.85667 | 0.886 | 0.88 | 0.823 | 0.85 | 0.85 | 0.86 | 0.88 | 0.88 | 0.8517 | 0.82691 | 0.779 | 0.85 | 0.86 | 0.85 | 0.84 | 0.78 | 0.81413 | 0.9241 |  |  |  | 1
| 1 | 2 | 0.8555 | 0.84929 | 0.85 | 0.91 | 0.82333 | 0.887 | 0.79 | 0.839 | 0.87 | 0.85 | 0.88 | 0.86 | 0.87 | 0.8038 | 0.82556 | 0.8 | 0.85 | 0.85 | 0.84 | 0.85 | 0.79 | 0.79853 | 0.9288 |  |  |  | 0
| 1 | 3 | 0.8542 | 0.84883 | 0.85 | 0.87 | 0.825 | 0.89 | 0.84 | 0.83 | 0.86 | 0.85 | 0.88 | 0.87 | 0.85 | 0.7915 | 0.85868 | 0.901 | 0.84 | 0.86 | 0.85 | 0.85 | 0.8 | 0.82457 | 0.9209 |  |  |  | 16
| 1 | 4 | 0.8547 | 0.85038 | 0.89 | 0.86 | 0.82667 | 0.884 | 0.94 | 0.843 | 0.86 | 0.88 | 0.86 | 0.87 | 0.88 | 0.816 | 0.84297 | 0.779 | 0.85 | 0.87 | 0.84 | 0.84 | 0.81 | 0.81329 | 0.9664 |  |  |  | 0.85542
| 1 | 5 | 0.8644 | 0.84347 | 0.87 | 0.89 | 0.86333 | 0.891 | 0.91 | 0.814 | 0.85 | 0.86 | 0.85 | 0.87 | 0.9 | 0.7946 | 0.83397 | 0.758 | 0.84 | 0.87 | 0.8 | 0.84 | 0.8 | 0.81677 | 0.9143 |  |  |  | 21
| 1 | 6 | 0.8616 | 0.85298 | 0.87 | 0.9 | 0.87 | 0.888 | 0.85 | 0.836 | 0.87 | 0.89 | 0.89 | 0.89 | 0.89 | 0.767 | 0.82549 | 0.779 | 0.83 | 0.86 | 0.84 | 0.85 | 0.79 | 0.799 | 0.8928
| Mean |  | 0.85725 | 0.8484 | 0.86167 | 0.88333 | 0.84417 | 0.88767 | 0.86833 | 0.83083 | 0.86 | 0.86333 | 0.87 | 0.87333 | 0.87833 | 0.8041 | 0.8356 | 0.79933 | 0.84333 | 0.86167 | 0.83667 | 0.845 | 0.795 | 0.81105 | 0.92455
| Median |  | 0.8551 | 0.84906 | 0.86 | 0.88 | 0.84167 | 0.8875 | 0.865 | 0.833 | 0.86 | 0.855 | 0.87 | 0.87 | 0.88 | 0.7992 | 0.83044 | 0.779 | 0.845 | 0.86 | 0.84 | 0.845 | 0.795 | 0.81371 | 0.9225
| Std Dev. |  | 0.0046073 | 0.0034368 | 0.018348 | 0.019664 | 0.021441 | 0.002582 | 0.053448 | 0.010834 | 0.0089443 | 0.017512 | 0.015492 | 0.010328 | 0.017224 | 0.0284 | 0.013176 | 0.051547 | 0.008165 | 0.0075277 | 0.018619 | 0.0054772 | 0.010488 | 0.010314 | 0.024067
| Rel.Std.Dev. |  | 0.5374% | 0.4051% | 2.129% | 2.226% | 2.54% | 0.2909% | 6.155% | 1.304% | 1.04% | 2.028% | 1.781% | 1.183% | 1.961% | 3.532% | 1.577% | 6.449% | 0.9682% | 0.8736% | 2.225% | 0.6482% | 1.319% | 1.272% | 2.603%
| PDM3 |  | 0.2141% | -0.8211% | 0.7304% | 3.263% | -1.315% | 3.77% | 1.51% | -2.874% | 0.5356% | 0.9252% | 1.705% | 2.094% | 2.679% | -5.999% | -2.317% | -6.556% | -1.413% | 0.7304% | -2.192% | -1.218% | -7.063% | -5.187% | 8.082%
| Z-Score (Absolute) |  | 0 | 0.37 | 0.19 | 1.09 | 0.55 | 1.28 | 0.46 | 1.11 | 0.12 | 0.26 | 0.53 | 0.67 | 0.88 | 2.23 | 0.91 | 2.43 | 0.58 | 0.19 | 0.86 | 0.51 | 2.61 | 1.94 | 2.82
| Analytical results for Cd in OREAS 238b (Certified Value 0.097 ppm) | [27 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.09 | 1.22 | < 0.5 | 0.4 | 0.83067 | 0.09 | 0.12 | 0.49 | 0.2 | 0.07 | 0.07 | 0.09 | 0.07 | 0.12 | 0.15206 | 0.1 | 0.112 | 0.11 | 9.29 | 0.15 | 0.07 | 0.07 | 0.1352 | < 5 | 0.2581 | 0.09782 | 1
| 1 | 2 | 0.07 | 1.21 | < 0.5 | 0.8 | 0.823 | 0.13 | 0.13 | 0.25 | 0.2 | 0.09 | 0.09 | 0.08 | 0.07 | 0.13 | 0.14901 | 0.1 | 0.121 | 0.08 | 9.02 | 0.08 | 0.07 | 0.06 | 0.1312 | < 5 | 0.2403 | 0.08366 | 30
| 1 | 3 | 0.09 | 1.17 | < 0.5 | 0.6 | 0.824 | 0.11 | 0.13 | 0.45 | 0.3 | 0.07 | 0.07 | 0.07 | 0.08 | 0.13 | 0.17088 | 0.1 | 0.139 | 0.07 | 9.62 | 0.1 | 0.07 | 0.07 | 0.1366 | < 5 | 0.2367 | 0.08415 | 16
| 1 | 4 | 0.09 | 1.18 | < 0.5 | 0.6 | 0.84 | 0.08 | 0.14 | 0.25 | 0.3 | 0.09 | 0.08 | 0.07 | 0.06 | 0.13 | 0.1326 | 0.1 | 0.12 | 0.09 | 9.26 | 0.02 | 0.07 | 0.09 | 0.1178 | < 5 | 0.2349 | 0.08015 | 0.097127
| 1 | 5 | 0.07 | 1.33 | < 0.5 | 0.8 | 0.82133 | 0.12 | 0.14 | 0.27 | 0.2 | 0.08 | 0.08 | 0.07 | 0.07 | 0.16 | 0.14965 | 0.1 | 0.117 | 0.09 | 9.24 | 0.08 | 0.08 | 0.08 | 0.1084 | < 5 | 0.2367 | 0.07969 | 22
| 1 | 6 | 0.09 | 1.23 | < 0.5 | 0.5 | 0.839 | 0.1 | 0.13 | 0.29 | 0.3 | 0.08 | 0.05 | 0.08 | 0.08 | 0.13 | 0.1375 | 0.1 | 0.125 | 0.08 | 9.21 | 0.12 | 0.08 | 0.09 | 0.1232 | < 5 | 0.2367 | 0.08742
| Mean |  | 0.083333 | 1.2233 |  | 0.61667 | 0.82967 | 0.105 | 0.13167 | 0.33333 | 0.25 | 0.08 | 0.073333 | 0.076667 | 0.071667 | 0.13333 | 0.14862 | 0.1 | 0.12233 | 0.086667 | 9.2733 | 0.091667 | 0.073333 | 0.076667 | 0.1254 |  | 0.24057 | 0.085482
| Median |  | 0.09 | 1.215 |  | 0.6 | 0.82733 | 0.105 | 0.13 | 0.28 | 0.25 | 0.08 | 0.075 | 0.075 | 0.07 | 0.13 | 0.14933 | 0.1 | 0.1205 | 0.085 | 9.25 | 0.09 | 0.07 | 0.075 | 0.1272 |  | 0.2367 | 0.083905
| Std Dev. |  | 0.010328 | 0.057155 |  | 0.16021 | 0.008257 | 0.018708 | 0.0075277 | 0.10764 | 0.054772 | 0.0089443 | 0.013663 | 0.008165 | 0.0075277 | 0.013663 | 0.013335 | 1.5202e-17 | 0.0092448 | 0.013663 | 0.1949 | 0.044008 | 0.005164 | 0.012111 | 0.011016 |  | 0.0087687 | 0.0066791
| Rel.Std.Dev. |  | 12.39% | 4.672% |  | 25.98% | 0.9952% | 17.82% | 5.717% | 32.29% | 21.91% | 11.18% | 18.63% | 10.65% | 10.5% | 10.25% | 8.973% | 1.52e-14% | 7.557% | 15.76% | 2.102% | 48.01% | 7.042% | 15.8% | 8.784% |  | 3.645% | 7.814%
| PDM3 |  | -14.2% | 1160% |  | 534.9% | 754.2% | 8.106% | 35.56% | 243.2% | 157.4% | -17.63% | -24.5% | -21.07% | -26.21% | 37.28% | 53.01% | 2.958% | 25.95% | -10.77% | 9448% | -5.622% | -24.5% | -21.07% | 29.11% |  | 147.7% | -11.99%
| Z-Score (Absolute) |  | 0.58 | 16.26 | 1.89 | NA | 10.44 | 0.26 | 0.14 | 3.12 | NA | 0.63 | 0.72 | 0.67 | 0.75 | 0.16 | 0.39 | NA | 0 | 0.53 | 135.12 | 0.45 | 0.72 | 0.67 | 0.05 | 35.11 | 1.75 | 0.54
Results from laboratories 1.06, 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 238b (Certified Value 81 ppm) | [27 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 0
1
| 1 | 1 | 82.93 | 89.0558 | 81.4 | 87.8 | 86.26 | 79 | 82.4 | 85.8 | 82.6 | 80 | 84.81 | 77.3824 | 66 | 77.26 | 75.32 | 80.22 | 84.82 | 77.44 | 68.6 | 78.8991 | 73.8214 | [5 empty] | 1
| 1 | 2 | 83.27 | 88.7824 | 82.7 | 89.9 | 86.54 | 80 | 82.4 | 88.4 | 84.3 | 78.7 | 76.76 | 79.0281 | 67 | 77.06 | 78.18 | 85.06 | 77.81 | 79.32 | 67.2 | 79.8871 | 70.5155 | [5 empty] | 31
| 1 | 3 | 83.79 | 91.4055 | 81.3 | 88.9 | 89.97 | 77 | 79.5 | 87.4 | 81.8 | 78.4 | 79.55 | 78.3827 | 70 | 78.95 | 76.83 | 85.11 | 77.44 | 79.95 | 70.3 | 81.5488 | 71.8345 | [5 empty] | 16
| 1 | 4 | 84.05 | 90.9813 | 81.8 | 88.9 | 93.27 | 77 | 81.7 | 86.9 | 83.9 | 79.5 | 80.08 | 77.7478 | 65 | 76.8 | 75.86 | 83.55 | 73.18 | 79.76 | 69.2 | 72.971 | 72.3534 | [5 empty] | 80.6391
| 1 | 5 | 83.94 | 87.7145 | 81.3 | 88.6 | 91.04 | 77 | 81.3 | 85.4 | 80.3 | 78.9 | 78.08 | 79.6638 | 66 | 80.79 | 77.51 | 81.76 | 81.98 | 76.9 | 69.7 | 76.3841 | 70.7867 | [5 empty] | 23
| 1 | 6 | 84.91 | 90.3183 | 84 | 88.6 | 90.19 | 76 | 81.3 | 79.4 | 85.1 | 79.7 | 75.11 | 77.2406 | 67 | 78.21 | 75.24 | 80.96 | 82.29 | 76.61 | 67.5 | 75.7554 | 71.3299
| Mean |  | 83.815 | 89.7096 | 82.0833 | 88.7833 | 89.545 | 77.6667 | 81.4333 | 85.55 | 83 | 79.2 | 79.065 | 78.2409 | 66.8333 | 78.1783 | 76.49 | 82.7767 | 79.5867 | 78.33 | 68.75 | 77.5743 | 71.7736
| Median |  | 83.865 | 89.687 | 81.6 | 88.75 | 90.08 | 77 | 81.5 | 86.35 | 83.25 | 79.2 | 78.815 | 78.0652 | 66.5 | 77.735 | 76.345 | 82.655 | 79.895 | 78.38 | 68.9 | 77.6416 | 71.5822
| Std Dev. |  | 0.68567 | 1.4235 | 1.0797 | 0.67946 | 2.7029 | 1.5055 | 1.069 | 3.2023 | 1.7776 | 0.6261 | 3.3525 | 0.96388 | 1.7224 | 1.5107 | 1.2116 | 2.1036 | 4.2245 | 1.5129 | 1.2243 | 3.1246 | 1.2072
| Rel.Std.Dev. |  | 0.8181% | 1.587% | 1.315% | 0.7653% | 3.018% | 1.938% | 1.313% | 3.743% | 2.142% | 0.7905% | 4.24% | 1.232% | 2.577% | 1.932% | 1.584% | 2.541% | 5.308% | 1.931% | 1.781% | 4.028% | 1.682%
| PDM3 |  | 3.938% | 11.25% | 1.791% | 10.1% | 11.04% | -3.686% | 0.9849% | 6.09% | 2.928% | -1.785% | -1.952% | -2.974% | -17.12% | -3.052% | -5.145% | 2.651% | -1.305% | -2.863% | -14.74% | -3.801% | -10.99%
| Z-Score (Absolute) |  | 1.08 | 2.46 | 0.67 | 2.24 | 2.42 | 0.36 | 0.52 | 1.49 | 0.89 | 0 | 0.03 | 0.22 | 2.89 | 0.24 | 0.63 | 0.84 | 0.09 | 0.2 | 2.44 | 0.38 | 1.74
| Analytical results for Co in OREAS 238b (Certified Value 17 ppm) | [27 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 16.9 | 15.8321 | 16 | 17 | 22.1633 | 17.5 | 17 | 17 | 16.1 | 18.4 | 17.5 | 17 | 17.1 | 17.6 | 17.399 | 17.9 | 14.9 | 16.6 | 22.73 | 17.1 | 17.3 | 17.4 | 19.57 | 20 | 20.479 | 16.5318 | 1
| 1 | 2 | 16.6 | 15.3234 | 16 | 17 | 21.7567 | 17.1 | 18 | 17 | 16.5 | 17.9 | 17.7 | 17.3 | 17.4 | 17.3 | 17.3562 | 16.6 | 14.63 | 17.7 | 22.22 | 16.4 | 15.8 | 17.7 | 18.715 | 20 | 19.8108 | 17.1349 | 32
| 1 | 3 | 17.1 | 15.4948 | 16 | 17 | 21.5033 | 17 | 18 | 17 | 17.3 | 17 | 17.9 | 16.7 | 17.6 | 17.5 | 16.0012 | 17 | 14.8 | 16.1 | 22.83 | 17.3 | 15.8 | 17.8 | 19.38 | 19 | 19.0091 | 16.9062 | 16
| 1 | 4 | 17.2 | 15.2677 | 17 | 17 | 21.7133 | 17 | 17 | 17 | 16.9 | 17.6 | 17.5 | 17.4 | 17.7 | 17 | 17.6212 | 16.7 | 15.15 | 16.9 | 22.56 | 16.8 | 14.9 | 17.8 | 18.81 | 21 | 19.5436 | 16.7978 | 17.0384
| 1 | 5 | 17.6 | 15.2033 | 17 | 17 | 22.49 | 17.5 | 16 | 15 | 17 | 16.9 | 17 | 16.4 | 17.4 | 18.9 | 16.5574 | 16.8 | 15.39 | 16.8 | 22.51 | 16.4 | 16.6 | 17.6 | 19.19 | 20 | 20.2117 | 16.3308 | 24
| 1 | 6 | 17.5 | 15.1712 | 17 | 17 | 22.3867 | 17.1 | 15 | 16 | 16.7 | 17.7 | 16.8 | 17.4 | 18.2 | 17.4 | 16.9769 | 17.2 | 15.01 | 16 | 22.73 | 16.2 | 16.2 | 17.9 | 18.24 | 20 | 20.3453 | 16.7639
| Mean |  | 17.15 | 15.3821 | 16.5 | 17 | 22.0022 | 17.2 | 16.8333 | 16.5 | 16.75 | 17.5833 | 17.4 | 17.0333 | 17.5667 | 17.6167 | 16.9853 | 17.0333 | 14.98 | 16.6833 | 22.5967 | 16.7 | 16.1 | 17.7 | 18.9842 | 20 | 19.8999 | 16.7442
| Median |  | 17.15 | 15.2956 | 16.5 | 17 | 21.96 | 17.1 | 17 | 17 | 16.8 | 17.65 | 17.5 | 17.15 | 17.5 | 17.45 | 17.1665 | 16.9 | 14.955 | 16.7 | 22.645 | 16.6 | 16 | 17.75 | 19 | 20 | 20.0113 | 16.7808
| Std Dev. |  | 0.37283 | 0.24829 | 0.54772 | 0 | 0.40108 | 0.23664 | 1.169 | 0.83666 | 0.41833 | 0.56362 | 0.41952 | 0.41312 | 0.37238 | 0.66156 | 0.61068 | 0.4761 | 0.26803 | 0.61779 | 0.21942 | 0.43818 | 0.81486 | 0.17889 | 0.48966 | 0.63246 | 0.55846 | 0.28218
| Rel.Std.Dev. |  | 2.174% | 1.614% | 3.32% | 0% | 1.823% | 1.376% | 6.945% | 5.071% | 2.497% | 3.205% | 2.411% | 2.425% | 2.12% | 3.755% | 3.595% | 2.795% | 1.789% | 3.703% | 0.971% | 2.624% | 5.061% | 1.011% | 2.579% | 3.162% | 2.806% | 1.685%
| PDM3 |  | 0.655% | -9.721% | -3.16% | -0.2253% | 29.13% | 0.9485% | -1.204% | -3.16% | -1.693% | 3.198% | 2.122% | -0.02971% | 3.1% | 3.394% | -0.3115% | -0.02971% | -12.08% | -2.084% | 32.62% | -1.986% | -5.508% | 3.883% | 11.42% | 17.38% | 16.79% | -1.727%
| Z-Score (Absolute) |  | 0.15 | 2.09 | 0.67 | 0.04 | 6.28 | 0.21 | 0.25 | 0.67 | 0.36 | 0.7 | 0.46 | 0 | 0.67 | 0.74 | 0.06 | 0 | 2.6 | 0.44 | 7.03 | 0.42 | 1.18 | 0.84 | 2.47 | 3.75 | 3.62 | 0.37
| Analytical results for Cr in OREAS 238b (Certified Value 124 ppm) | [27 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 148 | 110 | 122 | 128.1367 | 116 | 138 | 91 | 110 | 128 | 127 | 124 | 120 | 108 | 131.9587 | 137 | 59.18 | 136.1 | 121 | 116 | 115 | 131.4 | 119 | 97.2695 | 145.1231 |  |  | 1
| 1 | 2 | 147 | 110 | 154 | 126.5433 | 118 | 144 | 107 | 112 | 130 | 134 | 123 | 121 | 111 | 132.5985 | 140 | 96.61 | 134.7 | 122 | 115 | 111 | 129.2 | 120 | 98.1483 | 143.2511 |  |  | 33
| 1 | 3 | 151 | 110 | 109 | 126.0567 | 122 | 145 | 108 | 114 | 127 | 133 | 118 | 120 | 109 | 127.9708 | 151 | 111.3 | 135.9 | 124 | 121 | 117 | 132.9 | 121 | 99.6445 | 145.4305 |  |  | 16
| 1 | 4 | 150 | 110 | 116 | 125.6367 | 124 | 153 | 122 | 111 | 131 | 129 | 118 | 122 | 105 | 130.4177 | 136 | 120.7 | 136.9 | 132 | 124 | 117 | 132 | 121 | 95.0845 | 146.1387 |  |  | 123.9533
| 1 | 5 | 143 | 120 | 116 | 129.78 | 117 | 126 | 109 | 118 | 129 | 128 | 124 | 124 | 107 | 129.4675 | 132 | 124.3 | 139.9 | 125 | 127 | 113 | 132.3 | 124 | 95.8445 | 145.4961 |  |  | 25
| 1 | 6 | 153 | 120 | 117 | 128.1967 | 114 | 146 | 117 | 108 | 132 | 130 | 119 | 124 | 109 | 129.2151 | 136 | 124.9 | 133.5 | 132 | 120 | 113 | 129.9 | 125 | 95.6783 | 144.7164
| Mean |  | 148.6667 | 113.3333 | 122.3333 | 127.3917 | 118.5 | 142 | 109 | 112.1667 | 129.5 | 130.1667 | 121 | 121.8333 | 108.1667 | 130.2714 | 138.6667 | 106.165 | 136.1667 | 126 | 120.5 | 114.3333 | 131.2833 | 121.6667 | 96.9449 | 145.026
| Median |  | 149 | 110 | 116.5 | 127.34 | 117.5 | 144.5 | 108.5 | 111.5 | 129.5 | 129.5 | 121 | 121.5 | 108.5 | 129.9426 | 136.5 | 116 | 136 | 124.5 | 120.5 | 114 | 131.7 | 121 | 96.557 | 145.2768
| Std Dev. |  | 3.5024 | 5.164 | 16.0582 | 1.5804 | 3.7815 | 9.1869 | 10.6019 | 3.4881 | 1.8708 | 2.7869 | 2.9665 | 1.8348 | 2.0412 | 1.7515 | 6.5625 | 25.363 | 2.1824 | 4.858 | 4.5935 | 2.4221 | 1.4442 | 2.3381 | 1.7388 | 0.98775
| Rel.Std.Dev. |  | 2.356% | 4.556% | 13.13% | 1.241% | 3.191% | 6.47% | 9.727% | 3.11% | 1.445% | 2.141% | 2.452% | 1.506% | 1.887% | 1.345% | 4.733% | 23.89% | 1.603% | 3.856% | 3.812% | 2.118% | 1.1% | 1.922% | 1.794% | 0.6811%
| PDM3 |  | 19.94% | -8.568% | -1.307% | 2.774% | -4.399% | 14.56% | -12.06% | -9.509% | 4.475% | 5.013% | -2.383% | -1.71% | -12.74% | 5.097% | 11.87% | -14.35% | 9.853% | 1.651% | -2.786% | -7.761% | 5.914% | -1.845% | -21.79% | 17%
| Z-Score (Absolute) |  | 2.12 | 0.7 | 0.02 | 0.42 | 0.29 | 1.59 | 1.04 | 0.79 | 0.59 | 0.64 | 0.09 | 0.02 | 1.11 | 0.65 | 1.32 | 1.27 | 1.12 | 0.31 | 0.13 | 0.62 | 0.73 | 0.03 | 2 | 1.83
| Analytical results for Cs in OREAS 238b (Certified Value 9.52 ppm) | [27 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 9.69 | 8.4996 | 9.4 | 9.27 | 9.9 | 8.31 | 10.1 | 9.8 | 9.47 | 9.54 | 9.45 | 9.88 | 9.8 | 11 | 8.446 | 10 | 9.67 | 9.35 | 9.1 | 9.1629 | 9.3794 | [5 empty] | 1
| 1 | 2 | 9.73 | 8.2362 | 9.4 | 9.49 | 9.7 | 8.95 | 9.8 | 10.3 | 9.71 | 9.55 | 9.32 | 9.69 | 8.1 | 10.7 | 8.347 | 9 | 9.73 | 8.6 | 9.5 | 9.3304 | 9.5706 | [5 empty] | 34
| 1 | 3 | 9.81 | 8.4102 | 9.5 | 9.59 | 9.9 | 9.07 | 9.3 | 9.7 | 9.53 | 9.73 | 9.11 | 9.81 | 11.9 | 10.8 | 8.574 | 9 | 9.88 | 8.53 | 9.5 | 9.3937 | 9.3536 | [5 empty] | 16
| 1 | 4 | 9.77 | 8.334 | 9.8 | 9.48 | 9.7 | 9.62 | 9.6 | 10.6 | 9.89 | 9.37 | 9.54 | 9.91 | 9.3 | 10.5 | 8.321 | 9 | 9.69 | 8.09 | 9.7 | 9.4453 | 9.3469 | [5 empty] | 9.5152
| 1 | 5 | 9.73 | 8.1788 | 9.6 | 9.52 | 9.9 | 9.32 | 9.3 | 9.4 | 9.61 | 9.3 | 9.03 | 9.73 | 10.1 | 10.5 | 8.502 | 9 | 9.84 | 8.96 | 9.2 | 9.3843 | 9.1501 | [5 empty] | 26
| 1 | 6 | 9.73 | 8.3877 | 9.6 | 9.31 | 9.8 | 9.24 | 9.6 | 9 | 9.75 | 9.45 | 9.52 | 10.15 | 9.6 | 10.4 | 8.497 | 9 | 9.71 | 8.96 | 9 | 9.1314 | 9.5274
| Mean |  | 9.7433 | 8.3411 | 9.55 | 9.4433 | 9.8167 | 9.085 | 9.6167 | 9.8 | 9.66 | 9.49 | 9.3283 | 9.8617 | 9.8 | 10.65 | 8.4478 | 9.1667 | 9.7533 | 8.7483 | 9.3333 | 9.308 | 9.388
| Median |  | 9.73 | 8.3609 | 9.55 | 9.485 | 9.85 | 9.155 | 9.6 | 9.75 | 9.66 | 9.495 | 9.385 | 9.845 | 9.7 | 10.6 | 8.4715 | 9 | 9.72 | 8.78 | 9.35 | 9.3574 | 9.3665
| Std Dev. |  | 0.041312 | 0.11787 | 0.15166 | 0.12549 | 0.098319 | 0.44375 | 0.30605 | 0.5831 | 0.15427 | 0.15218 | 0.21591 | 0.16449 | 1.2394 | 0.22583 | 0.097495 | 0.40825 | 0.085946 | 0.43696 | 0.27325 | 0.1302 | 0.14985
| Rel.Std.Dev. |  | 0.424% | 1.413% | 1.588% | 1.329% | 1.002% | 4.884% | 3.182% | 5.95% | 1.597% | 1.604% | 2.315% | 1.668% | 12.65% | 2.12% | 1.154% | 4.454% | 0.8812% | 4.995% | 2.928% | 1.399% | 1.596%
| PDM3 |  | 2.398% | -12.34% | 0.3662% | -0.7548% | 3.169% | -4.521% | 1.067% | 2.994% | 1.522% | -0.2644% | -1.963% | 3.642% | 2.994% | 11.93% | -11.22% | -3.662% | 2.503% | -8.059% | -1.911% | -2.177% | -1.336%
| Z-Score (Absolute) |  | 0.66 | 3.48 | 0.09 | 0.23 | 0.88 | 1.28 | 0.29 | 0.83 | 0.41 | 0.09 | 0.57 | 1.01 | 0.83 | 3.34 | 3.17 | NA | 0.69 | 2.28 | 0.55 | 0.63 | 0.39
Results from laboratory 1.21 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 238b (Certified Value 30.8 ppm) | [27 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 1
2
| 1 | 1 | 32 | 27.7559 | 32 | 29 | 31.3867 | 33 | 27 | 27 | 28.8 | 31.2 | 31.7 | 32.3 | 30.8 | 30.3 | 30.9427 | 28.8 | 28.49 | 33.2 | 32.64 | 31.5 | 31.5 | 31 | 32.26 | 33 | 6 | 30.6165 | 1
| 1 | 2 | 32.1 | 29.0912 | 32 | 29 | 30.1833 | 31 | 34 | 29 | 29.9 | 31.4 | 32.3 | 32 | 30.5 | 29.9 | 32.0739 | 27.5 | 28.95 | 33.2 | 32.87 | 31.4 | 28.9 | 30.7 | 31.62 | 34 | 7 | 30.0256 | 35
| 1 | 3 | 33.2 | 28.2293 | 32 | 29 | 29.9533 | 33 | 30 | 26 | 29.8 | 29.5 | 31.9 | 31 | 30.8 | 30.2 | 32.3391 | 28.1 | 29.36 | 31 | 31.79 | 37.3 | 29 | 30.5 | 32.48 | 33 | 13 | 30.5813 | 16
| 1 | 4 | 32.8 | 29.3495 | 34 | 29 | 29.78 | 32 | 31 | 29 | 28.6 | 30.4 | 31.2 | 32.6 | 30.4 | 30.5 | 32.1842 | 30.8 | 29.14 | 33.9 | 30.02 | 32.8 | 29.2 | 30.6 | 31.96 | 34 | 6 | 30.6895 | 30.7977
| 1 | 5 | 32.6 | 25.9161 | 32 | 28 | 32.2533 | 31 | 30 | 28 | 28.3 | 29.7 | 31.4 | 31.2 | 31.2 | 30.1 | 31.7525 | 28 | 30.3 | 33 | 31.1 | 30.3 | 30.2 | 30.4 | 32.29 | 33 | 7 | 30.2497 | 27
| 1 | 6 | 32.6 | 28.8536 | 34 | 28 | 31.32 | 33 | 33 | 27 | 31.4 | 30 | 30.5 | 33.3 | 31.4 | 30.4 | 31.4138 | 28.8 | 29.49 | 31.2 | 32.81 | 30.8 | 30 | 31 | 31.87 | 33 | 7 | 30.4203
| Mean |  | 32.55 | 28.1993 | 32.6667 | 28.6667 | 30.8128 | 32.1667 | 30.8333 | 27.6667 | 29.4667 | 30.3667 | 31.5 | 32.0667 | 30.85 | 30.2333 | 31.7844 | 28.6667 | 29.2883 | 32.5833 | 31.8717 | 32.35 | 29.8 | 30.7 | 32.08 | 33.3333 | 7.6667 | 30.4305
| Median |  | 32.6 | 28.5415 | 32 | 29 | 30.7517 | 32.5 | 30.5 | 27.5 | 29.3 | 30.2 | 31.55 | 32.15 | 30.8 | 30.25 | 31.9132 | 28.45 | 29.25 | 33.1 | 32.215 | 31.45 | 29.6 | 30.65 | 32.11 | 33 | 7 | 30.5008
| Std Dev. |  | 0.44609 | 1.2615 | 1.0328 | 0.5164 | 0.98624 | 0.98319 | 2.4833 | 1.2111 | 1.1483 | 0.78655 | 0.6229 | 0.86641 | 0.38859 | 0.21602 | 0.52882 | 1.1587 | 0.60714 | 1.1907 | 1.14 | 2.5665 | 0.98995 | 0.25298 | 0.31831 | 0.5164 | 2.6583 | 0.25342
| Rel.Std.Dev. |  | 1.37% | 4.474% | 3.162% | 1.801% | 3.201% | 3.057% | 8.054% | 4.377% | 3.897% | 2.59% | 1.977% | 2.702% | 1.26% | 0.7145% | 1.664% | 4.042% | 2.073% | 3.654% | 3.577% | 7.934% | 3.322% | 0.824% | 0.9922% | 1.549% | 34.67% | 0.8328%
| PDM3 |  | 5.69% | -8.437% | 6.068% | -6.92% | 0.04882% | 4.445% | 0.1156% | -10.17% | -4.322% | -1.4% | 2.28% | 4.12% | 0.1697% | -1.833% | 3.204% | -6.92% | -4.901% | 5.798% | 3.487% | 5.04% | -3.24% | -0.3174% | 4.163% | 8.233% | -75.11% | -1.193%
| Z-Score (Absolute) |  | 0.9 | 1.36 | 0.96 | 1.12 | 0.01 | 0.7 | 0.01 | 1.64 | 0.7 | 0.24 | 0.35 | 0.64 | 0.01 | 0.31 | 0.5 | 1.12 | 0.8 | 0.91 | 0.54 | 0.79 | 0.53 | 0.06 | 0.65 | 1.3 | 12.01 | 0.2
| Analytical results for Dy in OREAS 238b (Certified Value 3.56 ppm) | [27 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS | 2
2
| 1 | 1 | 1.01: 3.1333; 1.03: 3.65; 1.05: 4.1; 1.06: 3.1; 1.07: 4.3; 1.08: 2.924; 1.09: 3.89; 1.10: 3; 1.11: 2.68; 1.12: 4.6928 | 1
| 1 | 2 | 1.01: 3.1988; 1.03: 3.7; 1.05: 4.1; 1.06: 2.7; 1.07: 4.29; 1.08: 2.926; 1.09: 3.9; 1.10: 3; 1.11: 2.71; 1.12: 4.4238 | 9
| 1 | 3 | 1.01: 3.1441; 1.03: 3.6; 1.05: 4.1; 1.06: 4.1; 1.07: 4.31; 1.08: 3; 1.09: 4.01; 1.10: 3; 1.11: 2.76; 1.12: 4.7196 | 16
| 1 | 4 | 1.01: 3.1188; 1.03: 3.9; 1.05: 4; 1.06: 2.9; 1.07: 4.16; 1.08: 3.069; 1.09: 3.7; 1.10: 3; 1.11: 2.67; 1.12: 4.4727 | 3.5604
| 1 | 5 | 1.01: 3.0547; 1.03: 3.95; 1.05: 4; 1.06: 3.4; 1.07: 4.24; 1.08: 2.896; 1.09: 3.94; 1.10: 3; 1.11: 2.71; 1.12: 4.4092 | 28
| 1 | 6 | 1.01: 3.1425; 1.03: 3.9; 1.05: 4.3; 1.06: 3.6; 1.07: 4.25; 1.08: 3.008; 1.09: 3.71; 1.10: 2.9; 1.11: 2.58; 1.12: 4.4837
| Mean |  | 1.01: 3.132; 1.03: 3.7833; 1.05: 4.1; 1.06: 3.3; 1.07: 4.2583; 1.08: 2.9705; 1.09: 3.8583; 1.10: 2.9833; 1.11: 2.685; 1.12: 4.5336
| Median |  | 1.01: 3.1379; 1.03: 3.8; 1.05: 4.1; 1.06: 3.25; 1.07: 4.27; 1.08: 2.963; 1.09: 3.895; 1.10: 3; 1.11: 2.695; 1.12: 4.4782
| Std Dev. |  | 1.01: 0.046643; 1.03: 0.15055; 1.05: 0.10954; 1.06: 0.5099; 1.07: 0.055648; 1.08: 0.065835; 1.09: 0.12608; 1.10: 0.040825; 1.11: 0.060249; 1.12: 0.13688
| Rel.Std.Dev. |  | 1.01: 1.489%; 1.03: 3.979%; 1.05: 2.672%; 1.06: 15.45%; 1.07: 1.307%; 1.08: 2.216%; 1.09: 3.268%; 1.10: 1.368%; 1.11: 2.244%; 1.12: 3.019%
| PDM3 |  | 1.01: -12.03%; 1.03: 6.26%; 1.05: 15.15%; 1.06: -7.315%; 1.07: 19.6%; 1.08: -16.57%; 1.09: 8.366%; 1.10: -16.21%; 1.11: -24.59%; 1.12: 27.33%
| Z-Score (Absolute) |  | 1.01: 0.49; 1.03: 0.29; 1.05: 0.67; 1.06: 0.29; 1.07: 0.87; 1.08: 0.69; 1.09: 0.38; 1.10: 0.67; 1.11: 1.03; 1.12: 1.2
| Analytical results for Er in OREAS 238b (Certified Value 1.9 ppm) | [27 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.6255; 1.03: 2; 1.05: 2.2; 1.06: 2.2; 1.07: 2.28; 1.08: 1.55; 1.09: 2.05; 1.10: 1.5; 1.11: 1.3303; 1.12: 2.6834 | 1
| 1 | 2 | 1.01: 1.5191; 1.03: 2; 1.05: 2.1; 1.06: 1.4; 1.07: 2.29; 1.08: 1.594; 1.09: 2.05; 1.10: 1.5; 1.11: 1.3575; 1.12: 2.5014 | 10
| 1 | 3 | 1.01: 1.6435; 1.03: 1.95; 1.05: 2; 1.06: 2; 1.07: 2.27; 1.08: 1.577; 1.09: 2.13; 1.10: 1.5; 1.11: 1.3753; 1.12: 2.6879 | 16
| 1 | 4 | 1.01: 1.6459; 1.03: 2.05; 1.05: 2.2; 1.06: 1.7; 1.07: 2.24; 1.08: 1.504; 1.09: 2.08; 1.10: 1.5; 1.11: 1.3665; 1.12: 2.5495 | 1.8998
| 1 | 5 | 1.01: 1.5862; 1.03: 2; 1.05: 2.2; 1.06: 2.6; 1.07: 2.3; 1.08: 1.528; 1.09: 2.1; 1.10: 1.5; 1.11: 1.3789; 1.12: 2.5271 | 29
| 1 | 6 | 1.01: 1.5807; 1.03: 2; 1.05: 2.1; 1.06: 1.6; 1.07: 2.3; 1.08: 1.539; 1.09: 2.09; 1.10: 1.5; 1.11: 1.3183; 1.12: 2.5362
| Mean |  | 1.01: 1.6002; 1.03: 2; 1.05: 2.1333; 1.06: 1.9167; 1.07: 2.28; 1.08: 1.5487; 1.09: 2.0833; 1.10: 1.5; 1.11: 1.3545; 1.12: 2.5809
| Median |  | 1.01: 1.6059; 1.03: 2; 1.05: 2.15; 1.06: 1.85; 1.07: 2.285; 1.08: 1.5445; 1.09: 2.085; 1.10: 1.5; 1.11: 1.362; 1.12: 2.5428
| Std Dev. |  | 1.01: 0.048492; 1.03: 0.031623; 1.05: 0.08165; 1.06: 0.44008; 1.07: 0.022804; 1.08: 0.032788; 1.09: 0.030768; 1.10: 0; 1.11: 0.024808; 1.12: 0.082637
| Rel.Std.Dev. |  | 1.01: 3.03%; 1.03: 1.581%; 1.05: 3.827%; 1.06: 22.96%; 1.07: 1%; 1.08: 2.117%; 1.09: 1.477%; 1.10: 0%; 1.11: 1.832%; 1.12: 3.202%
| PDM3 |  | 1.01: -15.77%; 1.03: 5.277%; 1.05: 12.3%; 1.06: 0.8902%; 1.07: 20.02%; 1.08: -18.48%; 1.09: 9.663%; 1.10: -21.04%; 1.11: -28.7%; 1.12: 35.85%
| Z-Score (Absolute) |  | 1.01: 0.71; 1.03: 0.08; 1.05: 0.35; 1.06: 0.08; 1.07: 0.64; 1.08: 0.81; 1.09: 0.25; 1.10: 0.91; 1.11: 1.2; 1.12: 1.24
| Analytical results for Eu in OREAS 238b (Certified Value 1.26 ppm) | [27 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS | 2
3
| 1 | 1 | 1.01: 1.6095; 1.03: 1.4; 1.05: 1.29; 1.06: 1.3; 1.07: 1.37; 1.08: 1.287; 1.09: 1.27; 1.10: 1.1; 1.11: 1.0328; 1.12: 1.168 | 1
| 1 | 2 | 1.01: 1.525; 1.03: 1.35; 1.05: 1.35; 1.06: 1; 1.07: 1.31; 1.08: 1.274; 1.09: 1.3; 1.10: 1.2; 1.11: 1.0826; 1.12: 1.0467 | 11
| 1 | 3 | 1.01: 1.5993; 1.03: 1.45; 1.05: 1.31; 1.06: 1.4; 1.07: 1.35; 1.08: 1.32; 1.09: 1.3; 1.10: 1.2; 1.11: 1.098; 1.12: 1.1339 | 16
| 1 | 4 | 1.01: 1.5174; 1.03: 1.4; 1.05: 1.33; 1.06: 1.1; 1.07: 1.32; 1.08: 1.273; 1.09: 1.26; 1.10: 1.2; 1.11: 1.0815; 1.12: 1.1229 | 1.2576
| 1 | 5 | 1.01: 1.6254; 1.03: 1.45; 1.05: 1.3; 1.06: 1.4; 1.07: 1.28; 1.08: 1.294; 1.09: 1.27; 1.10: 1.1; 1.11: 1.0951; 1.12: 1.0791 | 30
| 1 | 6 | 1.01: 1.6072; 1.03: 1.45; 1.05: 1.32; 1.06: 1.2; 1.07: 1.31; 1.08: 1.304; 1.09: 1.25; 1.10: 1.1; 1.11: 1.0422; 1.12: 1.0963
| Mean |  | 1.01: 1.5806; 1.03: 1.4167; 1.05: 1.3167; 1.06: 1.2333; 1.07: 1.3233; 1.08: 1.292; 1.09: 1.275; 1.10: 1.15; 1.11: 1.072; 1.12: 1.1078
| Median |  | 1.01: 1.6032; 1.03: 1.425; 1.05: 1.315; 1.06: 1.25; 1.07: 1.315; 1.08: 1.2905; 1.09: 1.27; 1.10: 1.15; 1.11: 1.082; 1.12: 1.1097
| Std Dev. |  | 1.01: 0.046868; 1.03: 0.040825; 1.05: 0.021602; 1.06: 0.1633; 1.07: 0.032042; 1.08: 0.018122; 1.09: 0.020736; 1.10: 0.054772; 1.11: 0.027702; 1.12: 0.042957
| Rel.Std.Dev. |  | 1.01: 2.965%; 1.03: 2.882%; 1.05: 1.641%; 1.06: 13.24%; 1.07: 2.421%; 1.08: 1.403%; 1.09: 1.626%; 1.10: 4.763%; 1.11: 2.584%; 1.12: 3.878%
| PDM3 |  | 1.01: 25.68%; 1.03: 12.64%; 1.05: 4.693%; 1.06: -1.933%; 1.07: 5.223%; 1.08: 2.732%; 1.09: 1.38%; 1.10: -8.559%; 1.11: -14.76%; 1.12: -11.91%
| Z-Score (Absolute) |  | 1.01: 2.63; 1.03: 1.07; 1.05: 0.12; 1.06: NA; 1.07: 0.18; 1.08: 0.12; 1.09: 0.28; 1.10: NA; 1.11: 2.21; 1.12: 1.87
Results from laboratories 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 238b (Certified Value 4.1 wt.%) | [27 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 4.11 | 4.0042 | 4.16 | 4.22 | 4.18 | 4.37 | 3.75 | 4.08 | 3.9 | 4.02 | 4.08 | 4.09 | 3.92 | 4.1524 | 4.1149 | 4.029 | 4.06 | 4.22 | 4.15 | 4.12 | 3.99 | 4.0371 | 4.216 | 4.1435 |  |  | 1
| 1 | 2 | 4.04 | 4.0441 | 4.17 | 4.45 | 4.1567 | 4.36 | 3.9 | 4.17 | 3.9 | 4.07 | 4.19 | 4.01 | 3.88 | 3.9964 | 4.2218 | 4.134 | 4.09 | 4.17 | 4.18 | 4.15 | 4 | 3.9239 | 4.2623 | 4.2045 |  |  | 0
| 1 | 3 | 4.16 | 4.0533 | 4.16 | 4.21 | 4.17 | 4.39 | 3.98 | 4.1 | 3.83 | 4.07 | 4.2 | 4.03 | 3.8 | 4.029 | 4.1353 | 4.644 | 4.02 | 4.19 | 4.17 | 4.16 | 4.02 | 4.0274 | 4.1781 | 4.2554 |  |  | 16
| 1 | 4 | 4.22 | 4.0498 | 4.16 | 4.27 | 4.195 | 4.37 | 4.03 | 4.17 | 3.82 | 4.15 | 4.09 | 4.04 | 3.92 | 4.082 | 4.1239 | 4.001 | 4.1 | 4.18 | 4.18 | 4.15 | 4 | 3.9992 | 4.3865 | 4.0598 |  |  | 4.1016
| 1 | 5 | 4.22 | 4.0341 | 4.19 | 4.38 | 4.23 | 4.38 | 3.99 | 4 | 3.81 | 4.1 | 4.04 | 4.03 | 4.01 | 4.0351 | 4.1858 | 3.91 | 4.1 | 4.19 | 4.03 | 4.14 | 4.01 | 4.0499 | 4.1964 | 4.1205 |  |  | 31
| 1 | 6 | 4.23 | 4.0102 | 4.1 | 4.51 | 4.175 | 4.37 | 3.89 | 4.15 | 3.81 | 4.23 | 4.14 | 4.11 | 3.97 | 3.9688 | 4.226 | 4.022 | 4.01 | 4.18 | 4.15 | 4.16 | 3.98 | 3.9626 | 4.0699 | 4.0985
| Mean |  | 4.1633 | 4.0326 | 4.1567 | 4.34 | 4.1844 | 4.3733 | 3.9233 | 4.1117 | 3.845 | 4.1067 | 4.1233 | 4.0517 | 3.9167 | 4.0439 | 4.1679 | 4.1233 | 4.0633 | 4.1883 | 4.1433 | 4.1467 | 4 | 4 | 4.2182 | 4.147
| Median |  | 4.19 | 4.0391 | 4.16 | 4.325 | 4.1775 | 4.37 | 3.94 | 4.125 | 3.825 | 4.085 | 4.115 | 4.035 | 3.92 | 4.032 | 4.1606 | 4.0255 | 4.075 | 4.185 | 4.16 | 4.15 | 4 | 4.0133 | 4.2062 | 4.132
| Std Dev. |  | 0.07607 | 0.020817 | 0.030111 | 0.12554 | 0.025597 | 0.010328 | 0.10073 | 0.066156 | 0.043243 | 0.073937 | 0.064083 | 0.0392 | 0.072847 | 0.065452 | 0.049816 | 0.2649 | 0.040332 | 0.017224 | 0.057155 | 0.015055 | 0.014142 | 0.048596 | 0.10426 | 0.071732
| Rel.Std.Dev. |  | 1.827% | 0.5162% | 0.7244% | 2.893% | 0.6117% | 0.2362% | 2.567% | 1.609% | 1.125% | 1.8% | 1.554% | 0.9675% | 1.86% | 1.619% | 1.195% | 6.424% | 0.9926% | 0.4112% | 1.379% | 0.3631% | 0.3536% | 1.215% | 2.472% | 1.73%
| PDM3 |  | 1.505% | -1.683% | 1.342% | 5.812% | 2.019% | 6.625% | -4.347% | 0.2451% | -6.256% | 0.1232% | 0.5295% | -1.218% | -4.509% | -1.406% | 1.617% | 0.5295% | -0.9333% | 2.114% | 1.017% | 1.098% | -2.477% | -2.477% | 2.842% | 1.107%
| Z-Score (Absolute) |  | 0.43 | 0.97 | 0.36 | 2.32 | 0.65 | 2.67 | 2.14 | 0.12 | 2.97 | 0.18 | 0 | 0.77 | 2.21 | 0.85 | 0.48 | 0 | 0.64 | 0.7 | 0.21 | 0.25 | 1.32 | 1.32 | 1.01 | 0.25
| Analytical results for Ga in OREAS 238b (Certified Value 19.6 ppm) | [27 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 20.01 | 15.1206 | 19.2 | 21.3 | 22.94 | 19.5 | 18.29 | 21.8 | 20 | 19.4 | 21.1 | 20.3 | 19.5 | 19.35 | 18.7122 | 18 | 20.4 | 18.81 | 20.44 | 18.11 | 20.3893 | 24 | 19.1941 |  |  |  | 1
| 1 | 2 | 19.7 | 15.2411 | 19.2 | 22.5 | 21.7733 | 19.3 | 19.63 | 21.9 | 21 | 19.8 | 21.1 | 20.2 | 19.2 | 18.26 | 19.0748 | 17.2 | 20.9 | 18.94 | 17.84 | 18.08 | 19.9241 | 24 | 18.9264 |  |  |  | 39
| 1 | 3 | 19.97 | 15.6273 | 19.2 | 22.1 | 22.3167 | 19.3 | 19.78 | 22 | 21 | 18.4 | 21.5 | 19.25 | 19.3 | 17.94 | 19.2124 | 17.3 | 19.9 | 19.01 | 17.65 | 18.19 | 19.5365 | 24 | 19.2866 |  |  |  | 16
| 1 | 4 | 20.07 | 15.5884 | 19.8 | 21.8 | 22.4467 | 19.5 | 20.38 | 20.8 | 20 | 18.95 | 21.4 | 20.6 | 19 | 18.1 | 18.6233 | 17.1 | 19.6 | 18.8 | 16.74 | 18.17 | 19.3932 | 26 | 19.423 |  |  |  | 19.5512
| 1 | 5 | 20.14 | 15.3392 | 19.4 | 22.4 | 23.37 | 19.5 | 20.12 | 18.6 | 21 | 18.6 | 21.2 | 19.35 | 19.8 | 19.1 | 19.6867 | 17.1 | 20.6 | 18.93 | 18.89 | 18.14 | 19.4919 | 23 | 19.2105 |  |  |  | 32
| 1 | 6 | 19.89 | 15.0545 | 19.8 | 22.5 | 22.3067 | 19.4 | 19.51 | 19.9 | 21 | 18.9 | 19.15 | 20.8 | 20 | 18.56 | 18.8468 | 17.2 | 19.8 | 18.59 | 18.57 | 18.11 | 19.0905 | 26 | 19.4208
| Mean |  | 19.9633 | 15.3285 | 19.4333 | 22.1 | 22.5256 | 19.4167 | 19.6183 | 20.8333 | 20.6667 | 19.0083 | 20.9083 | 20.0833 | 19.4667 | 18.5517 | 19.0261 | 17.3167 | 20.2 | 18.8467 | 18.355 | 18.1333 | 19.6376 | 24.5 | 19.2436
| Median |  | 19.99 | 15.2901 | 19.3 | 22.25 | 22.3817 | 19.45 | 19.705 | 21.3 | 21 | 18.925 | 21.15 | 20.25 | 19.4 | 18.41 | 18.9608 | 17.2 | 20.15 | 18.87 | 18.205 | 18.125 | 19.5142 | 24 | 19.2486
| Std Dev. |  | 0.15462 | 0.23788 | 0.29439 | 0.47749 | 0.5565 | 0.098319 | 0.72585 | 1.3633 | 0.5164 | 0.51616 | 0.87659 | 0.64395 | 0.37771 | 0.56591 | 0.39153 | 0.34303 | 0.5099 | 0.14949 | 1.2689 | 0.041312 | 0.45554 | 1.2247 | 0.18417
| Rel.Std.Dev. |  | 0.7745% | 1.552% | 1.515% | 2.161% | 2.471% | 0.5064% | 3.7% | 6.544% | 2.499% | 2.715% | 4.193% | 3.206% | 1.94% | 3.05% | 2.058% | 1.981% | 2.524% | 0.7932% | 6.913% | 0.2278% | 2.32% | 4.999% | 0.9571%
| PDM3 |  | 2.108% | -21.6% | -0.6028% | 13.04% | 15.21% | -0.6881% | 0.3434% | 6.558% | 5.705% | -2.777% | 6.941% | 2.722% | -0.4323% | -5.112% | -2.686% | -11.43% | 3.318% | -3.604% | -6.118% | -7.252% | 0.4419% | 25.31% | -1.573%
| Z-Score (Absolute) |  | 0.46 | 3.81 | 0.03 | 2.42 | 2.81 | 0.05 | 0.14 | 1.26 | 1.1 | 0.42 | 1.33 | 0.57 | 0 | 0.84 | 0.41 | 1.98 | 0.67 | 0.57 | 1.02 | 1.23 | 0.16 | 4.63 | 0.21
| Analytical results for Gd in OREAS 238b (Certified Value 5.14 ppm) | [27 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS |  |  | 2
3
| 1 | 1 | 1.01: 6.9662; 1.03: 5.4; 1.05: 5.3; 1.06: 5.2; 1.07: 5.3; 1.08: 5.158; 1.09: 5.24; 1.10: 4.4; 1.11: 4.7641; 1.12: 0.77514 |  |  | 1
| 1 | 2 | 1.01: 6.9248; 1.03: 5.8; 1.05: 5.5; 1.06: 4.1; 1.07: 5; 1.08: 5.135; 1.09: 5.37; 1.10: 4.5; 1.11: 4.7624; 1.12: 0.73267 |  |  | 12
| 1 | 3 | 1.01: 6.9322; 1.03: 5.8; 1.05: 5.6; 1.06: 5.4; 1.07: 5.2; 1.08: 5.249; 1.09: 5.31; 1.10: 4.5; 1.11: 4.8825; 1.12: 0.77664 |  |  | 16
| 1 | 4 | 1.01: 6.9907; 1.03: 5.8; 1.05: 5.5; 1.06: 5; 1.07: 5; 1.08: 5.363; 1.09: 5.29; 1.10: 4.6; 1.11: 4.8066; 1.12: 0.74017 |  |  | 5.1398
| 1 | 5 | 1.01: 7.011; 1.03: 5.4; 1.05: 5.5; 1.06: 5.1; 1.07: 5.1; 1.08: 5.282; 1.09: 5.17; 1.10: 4.5; 1.11: 4.8515; 1.12: 0.72512 |  |  | 33
| 1 | 6 | 1.01: 6.899; 1.03: 5.6; 1.05: 5.6; 1.06: 5; 1.07: 5.1; 1.08: 5.269; 1.09: 5.01; 1.10: 4.3; 1.11: 4.6572; 1.12: 0.74553
| Mean |  | 1.01: 6.954; 1.03: 5.6333; 1.05: 5.5; 1.06: 4.9667; 1.07: 5.1167; 1.08: 5.2427; 1.09: 5.2317; 1.10: 4.4667; 1.11: 4.7874; 1.12: 0.74921
| Median |  | 1.01: 6.9492; 1.03: 5.7; 1.05: 5.5; 1.06: 5.05; 1.07: 5.1; 1.08: 5.259; 1.09: 5.265; 1.10: 4.5; 1.11: 4.7854; 1.12: 0.74285
| Std Dev. |  | 1.01: 0.04264; 1.03: 0.19664; 1.05: 0.10954; 1.06: 0.45019; 1.07: 0.1169; 1.08: 0.084287; 1.09: 0.12781; 1.10: 0.10328; 1.11: 0.07955; 1.12: 0.02179
| Rel.Std.Dev. |  | 1.01: 0.6132%; 1.03: 3.491%; 1.05: 1.992%; 1.06: 9.064%; 1.07: 2.285%; 1.08: 1.608%; 1.09: 2.443%; 1.10: 2.312%; 1.11: 1.662%; 1.12: 2.908%
| PDM3 |  | 1.01: 35.3%; 1.03: 9.602%; 1.05: 7.008%; 1.06: -3.368%; 1.07: -0.45%; 1.08: 2.001%; 1.09: 1.787%; 1.10: -13.1%; 1.11: -6.857%; 1.12: -85.42%
| Z-Score (Absolute) |  | 1.01: 3.37; 1.03: 0.87; 1.05: 0.62; 1.06: 0.39; 1.07: 0.11; 1.08: 0.13; 1.09: 0.11; 1.10: 1.34; 1.11: 0.73; 1.12: 8.37
| Analytical results for Ge in OREAS 238b (Indicative Value 0.2 ppm) | [27 empty] | Indicative
Ge | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.03: 1.06; 1.05: 1.08; 1.06: 1.09; 1.07: 1.10; 1.08: 1.11; 1.09: 1.12; 1.10: 1.13; 1.11: 1.14; 1.12: 1.15; 1.13: 1.16; 1.14: 1.26; 1.15: 1.27; 1.16: 1.30 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*OES/MS; 1.05: 4A*OES; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*OES/MS; 1.11: 4A*OES/MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*MS; 1.15: 4A*OES/MS; 1.16: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 1.8; 1.03: 0.2; 1.05: 1.2; 1.06: 0.18; 1.07: 2.1; 1.08: 3.9; 1.09: 0.17; 1.10: 0.25; 1.11: 0.08; 1.12: 0.19; 1.13: 0.44; 1.14: < 0.05; 1.15: 0.15; 1.16: 0.5105 |  |  | 1
| 1 | 2 | 1.01: 1.6; 1.03: 0.1; 1.05: 1.3; 1.06: 0.19; 1.07: 1.8; 1.08: 4.1; 1.09: 0.16; 1.10: 0.27; 1.11: 0.09; 1.12: 0.19; 1.13: 0.33; 1.14: < 0.05; 1.15: 0.13; 1.16: 0.2394 |  |  | 5
| 1 | 3 | 1.01: 1.8; 1.03: 0.3; 1.05: 1.3; 1.06: 0.2; 1.07: 1.6; 1.08: 4.2; 1.09: 0.15; 1.10: 0.24; 1.11: 0.09; 1.12: 0.2; 1.13: 0.49; 1.14: < 0.05; 1.15: 0.11; 1.16: 0.1877 |  |  | 16
| 1 | 4 | 1.01: 1.8; 1.03: 0.3; 1.05: 1.3; 1.06: 0.2; 1.07: 1.4; 1.08: 3.9; 1.09: 0.14; 1.10: 0.27; 1.11: 0.09; 1.12: 0.18; 1.13: 0.33; 1.14: < 0.05; 1.15: 0.1; 1.16: 0.1936 |  |  | 0.19974
| 1 | 5 | 1.01: 1.6; 1.03: 0.2; 1.05: 1.3; 1.06: 0.2; 1.07: 1.2; 1.08: 3.7; 1.09: 0.16; 1.10: 0.25; 1.11: 0.08; 1.12: 0.18; 1.13: 0.36; 1.14: < 0.05; 1.15: 0.11; 1.16: 0.1892 |  |  | 11
| 1 | 6 | 1.01: 1.8; 1.03: 0.2; 1.05: 1.3; 1.06: 0.21; 1.07: 1.5; 1.08: 4; 1.09: 0.15; 1.10: 0.12; 1.11: 0.11; 1.12: 0.19; 1.13: 0.52; 1.14: < 0.05; 1.15: 0.11; 1.16: 0.1549
| Mean |  | 1.01: 1.7333; 1.03: 0.21667; 1.05: 1.2833; 1.06: 0.19667; 1.07: 1.6; 1.08: 3.9667; 1.09: 0.155; 1.10: 0.23333; 1.11: 0.09; 1.12: 0.18833; 1.13: 0.41167; 1.15: 0.11833; 1.16: 0.24588
| Median |  | 1.01: 1.8; 1.03: 0.2; 1.05: 1.3; 1.06: 0.2; 1.07: 1.55; 1.08: 3.95; 1.09: 0.155; 1.10: 0.25; 1.11: 0.09; 1.12: 0.19; 1.13: 0.4; 1.15: 0.11; 1.16: 0.1914
| Std Dev. |  | 1.01: 0.10328; 1.03: 0.075277; 1.05: 0.040825; 1.06: 0.010328; 1.07: 0.31623; 1.08: 0.17512; 1.09: 0.010488; 1.10: 0.056804; 1.11: 0.010954; 1.12: 0.0075277; 1.13: 0.083287; 1.15: 0.018348; 1.16: 0.13242
| Rel.Std.Dev. |  | 1.01: 5.958%; 1.03: 34.74%; 1.05: 3.181%; 1.06: 5.252%; 1.07: 19.76%; 1.08: 4.415%; 1.09: 6.767%; 1.10: 24.34%; 1.11: 12.17%; 1.12: 3.997%; 1.13: 20.23%; 1.15: 15.51%; 1.16: 53.85%
| PDM3 |  | 1.01: 767.8%; 1.03: 8.476%; 1.05: 542.5%; 1.06: -1.537%; 1.07: 701.1%; 1.08: 1886%; 1.09: -22.4%; 1.10: 16.82%; 1.11: -54.94%; 1.12: -5.709%; 1.13: 106.1%; 1.15: -40.76%; 1.16: 23.1%
| Z-Score (Absolute) |  | 1.01: 8.42; 1.03: 0.05; 1.05: 5.91; 1.06: 0.16; 1.07: 7.67; 1.08: 20.88; 1.09: 0.39; 1.10: 0.05; 1.11: 0.75; 1.12: 0.2; 1.13: 1.04; 1.14: 1.12; 1.15: 0.6; 1.16: 0.12
| Analytical results for Hf in OREAS 238b (Certified Value 3.89 ppm) | [27 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 4.07 | 3.8214 | 8.4 | 4.5 | 3.8 | 3.3 | 3.9 | 3.7 | 3.9 | 3.9 | 3.8 | 3.8 | 3.65 | 4.4 | 3.41 | 4.54 | 3.9 | 3.19 | 3.0189 | 2.555 | 6.1723 | [5 empty] | 1
| 1 | 2 | 4 | 3.9064 | 8.4 | 4.8 | 3.8 | 3.6 | 4 | 3.8 | 4 | 4.1 | 4.2 | 3.6 | 3.51 | 4.3 | 3.69 | 4.15 | 3.6 | 3.29 | 3.092 | 2.6715 | 5.877 | [5 empty] | 25
| 1 | 3 | 4.09 | 3.7666 | 8.2 | 4.3 | 4 | 3.7 | 4 | 3.9 | 3.8 | 4.1 | 3.9 | 3.6 | 3.91 | 4.3 | 3.55 | 4.27 | 3.6 | 3.28 | 3.0675 | 2.7629 | 6.1319 | [5 empty] | 16
| 1 | 4 | 4.01 | 3.8011 | 8.4 | 4.4 | 3.8 | 3.8 | 3.9 | 3.8 | 3.9 | 4 | 3.8 | 3.7 | 3.74 | 4.2 | 3.45 | 4.35 | 3.5 | 3.31 | 3.0873 | 2.264 | 5.8998 | [5 empty] | 3.8854
| 1 | 5 | 3.98 | 3.7778 | 8.2 | 4.5 | 3.8 | 3.8 | 4 | 3.8 | 3.8 | 3.8 | 3.7 | 3.7 | 3.9 | 4.3 | 3.67 | 4.13 | 3.8 | 3.22 | 3.0691 | 2.1892 | 6.1568 | [5 empty] | 34
| 1 | 6 | 3.94 | 3.8926 | 7.8 | 4.6 | 4 | 3.8 | 3.9 | 4.5 | 4 | 3.8 | 3.9 | 3.8 | 3.9 | 4.3 | 3.5 | 4.57 | 3.8 | 3.18 | 3.0075 | 2.1726 | 5.9503
| Mean |  | 4.015 | 3.8277 | 8.2333 | 4.5167 | 3.8667 | 3.6667 | 3.95 | 3.9167 | 3.9 | 3.95 | 3.8833 | 3.7 | 3.7683 | 4.3 | 3.545 | 4.335 | 3.7 | 3.245 | 3.057 | 2.4359 | 6.0314
| Median |  | 4.005 | 3.8112 | 8.3 | 4.5 | 3.8 | 3.75 | 3.95 | 3.8 | 3.9 | 3.95 | 3.85 | 3.7 | 3.82 | 4.3 | 3.525 | 4.31 | 3.7 | 3.25 | 3.0683 | 2.4095 | 6.0411
| Std Dev. |  | 0.056125 | 0.058942 | 0.23381 | 0.17224 | 0.10328 | 0.19664 | 0.054772 | 0.29269 | 0.089443 | 0.13784 | 0.17224 | 0.089443 | 0.1651 | 0.063246 | 0.11485 | 0.18865 | 0.15492 | 0.055408 | 0.035501 | 0.25937 | 0.13669
| Rel.Std.Dev. |  | 1.398% | 1.54% | 2.84% | 3.813% | 2.671% | 5.363% | 1.387% | 7.473% | 2.293% | 3.49% | 4.435% | 2.417% | 4.381% | 1.471% | 3.24% | 4.352% | 4.187% | 1.707% | 1.161% | 10.65% | 2.266%
| PDM3 |  | 3.336% | -1.486% | 111.9% | 16.25% | -0.4821% | -5.63% | 1.663% | 0.8047% | 0.3758% | 1.663% | -0.05316% | -4.772% | -3.013% | 10.67% | -8.761% | 11.57% | -4.772% | -16.48% | -21.32% | -37.31% | 55.23%
| Z-Score (Absolute) |  | 0.48 | 0.2 | 16 | 2.33 | 0.06 | 0.8 | 0.25 | 0.12 | 0.06 | 0.25 | 0 | 0.67 | 0.42 | 1.53 | 1.24 | 1.66 | 0.67 | 2.35 | 3.04 | 5.32 | 7.9
| Analytical results for Hg in OREAS 238b (Indicative Value 0.03 ppm) | [27 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.03: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.03: 1.30 |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.03: 4A*MS |  |  | 3
3
| 1 | 1 | 1.01: < 2; 1.03: 0.0236 |  |  | 1
| 1 | 2 | 1.01: < 2; 1.03: 0.0332 |  |  | 6
| 1 | 3 | 1.01: < 2; 1.03: 0.0289 |  |  | 16
| 1 | 4 | 1.01: < 2; 1.03: 0.0356 |  |  | 0.029867
| 1 | 5 | 1.01: < 2; 1.03: 0.0278 |  |  | 12
| 1 | 6 | 1.01: < 2; 1.03: 0.0301
| Mean |  | 1.03: 0.029867
| Median |  | 1.03: 0.0295
| Std Dev. |  | 1.03: 0.0042075
| Rel.Std.Dev. |  | 1.03: 14.09%
| PDM3 |  | 1.03: -1.332e-13%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.67
| Analytical results for Ho in OREAS 238b (Certified Value 0.66 ppm) | [27 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.57636; 1.03: 0.7; 1.05: 0.8; 1.06: 0.6; 1.07: 0.76; 1.08: 0.531; 1.09: 0.74; 1.10: 0.5; 1.11: 0.4725; 1.12: 0.9428 |  |  | 1
| 1 | 2 | 1.01: 0.55429; 1.03: 0.7; 1.05: 0.7; 1.06: 0.5; 1.07: 0.74; 1.08: 0.537; 1.09: 0.76; 1.10: 0.5; 1.11: 0.4838; 1.12: 0.86474 |  |  | 14
| 1 | 3 | 1.01: 0.59606; 1.03: 0.68; 1.05: 0.8; 1.06: 0.9; 1.07: 0.76; 1.08: 0.54; 1.09: 0.74; 1.10: 0.5; 1.11: 0.4858; 1.12: 0.92151 |  |  | 16
| 1 | 4 | 1.01: 0.57046; 1.03: 0.7; 1.05: 0.8; 1.06: 0.6; 1.07: 0.74; 1.08: 0.608; 1.09: 0.75; 1.10: 0.6; 1.11: 0.4798; 1.12: 0.89046 |  |  | 0.65576
| 1 | 5 | 1.01: 0.57322; 1.03: 0.68; 1.05: 0.8; 1.06: 0.6; 1.07: 0.74; 1.08: 0.55; 1.09: 0.75; 1.10: 0.5; 1.11: 0.4874; 1.12: 0.86742 |  |  | 35
| 1 | 6 | 1.01: 0.59608; 1.03: 0.7; 1.05: 0.8; 1.06: 0.6; 1.07: 0.74; 1.08: 0.538; 1.09: 0.72; 1.10: 0.5; 1.11: 0.4738; 1.12: 0.89383
| Mean |  | 1.01: 0.57774; 1.03: 0.69333; 1.05: 0.78333; 1.06: 0.63333; 1.07: 0.74667; 1.08: 0.55067; 1.09: 0.74333; 1.10: 0.51667; 1.11: 0.48052; 1.12: 0.89679
| Median |  | 1.01: 0.57479; 1.03: 0.7; 1.05: 0.8; 1.06: 0.6; 1.07: 0.74; 1.08: 0.539; 1.09: 0.745; 1.10: 0.5; 1.11: 0.4818; 1.12: 0.89214
| Std Dev. |  | 1.01: 0.016105; 1.03: 0.010328; 1.05: 0.040825; 1.06: 0.13663; 1.07: 0.010328; 1.08: 0.028759; 1.09: 0.013663; 1.10: 0.040825; 1.11: 0.0062605; 1.12: 0.03055
| Rel.Std.Dev. |  | 1.01: 2.788%; 1.03: 1.49%; 1.05: 5.212%; 1.06: 21.57%; 1.07: 1.383%; 1.08: 5.223%; 1.09: 1.838%; 1.10: 7.902%; 1.11: 1.303%; 1.12: 3.407%
| PDM3 |  | 1.01: -11.9%; 1.03: 5.73%; 1.05: 19.45%; 1.06: -3.42%; 1.07: 13.86%; 1.08: -16.03%; 1.09: 13.35%; 1.10: -21.21%; 1.11: -26.72%; 1.12: 36.76%
| Z-Score (Absolute) |  | 1.01: 0.58; 1.03: 0.2; 1.05: 0.82; 1.06: 0.2; 1.07: 0.57; 1.08: 0.77; 1.09: 0.54; 1.10: 1; 1.11: 1.24; 1.12: 1.59
| Analytical results for In in OREAS 238b (Certified Value 0.068 ppm) | [27 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [7 empty] | 3
3
| 1 | 1 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.05 | 0.041 | 0.14 | 0.07 | 0.074 | 0.068 | 0.066 | 0.069 | 0.07 | 0.07 | 0.062 | 0.066 | 0.06 | 0.0769 | 0.06835 | [7 empty] | 1
| 1 | 2 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.08 | 0.04 | 0.12 | 0.08 | 0.071 | 0.075 | 0.066 | 0.067 | 0.07 | 0.07 | 0.082 | 0.056 | 0.06 | 0.0807 | 0.0656 | [7 empty] | 27
| 1 | 3 | 0.07 | < 0.05 | < 0.1 | < 2 | 0.07 | 0.043 | 0.13 | 0.08 | 0.073 | 0.074 | 0.064 | 0.066 | 0.07 | 0.07 | 0.086 | 0.057 | 0.06 | 0.0853 | 0.06574 | [7 empty] | 16
| 1 | 4 | 0.07 | 0.1 | < 0.1 | < 2 | 0.06 | 0.043 | 0.09 | 0.09 | 0.076 | 0.066 | 0.07 | 0.064 | 0.07 | 0.07 | 0.074 | 0.054 | 0.06 | 0.0828 | 0.06767 | [7 empty] | 0.068319
| 1 | 5 | 0.07 | 0.05 | < 0.1 | < 2 | 0.08 | 0.041 | 0.07 | 0.08 | 0.071 | 0.073 | 0.067 | 0.068 | 0.07 | 0.07 | 0.076 | 0.061 | 0.06 | 0.079 | 0.06604 | [7 empty] | 36
| 1 | 6 | 0.07 | 0.05 | < 0.1 | < 2 | 0.06 | 0.047 | 0.07 | 0.08 | 0.073 | 0.075 | 0.065 | 0.068 | 0.07 | 0.07 | 0.064 | 0.06 | 0.06 | 0.0745 | 0.06915
| Mean |  | 0.07 | 0.066667 |  |  | 0.066667 | 0.0425 | 0.10333 | 0.08 | 0.073 | 0.071833 | 0.066333 | 0.067 | 0.07 | 0.07 | 0.074 | 0.059 | 0.06 | 0.079867 | 0.067092
| Median |  | 0.07 | 0.05 |  |  | 0.065 | 0.042 | 0.105 | 0.08 | 0.073 | 0.0735 | 0.066 | 0.0675 | 0.07 | 0.07 | 0.075 | 0.0585 | 0.06 | 0.07985 | 0.066855
| Std Dev. |  | 0 | 0.028868 |  |  | 0.012111 | 0.00251 | 0.030768 | 0.0063246 | 0.0018974 | 0.0038687 | 0.0020656 | 0.0017889 | 0 | 0 | 0.0095499 | 0.0042895 | 0 | 0.0039287 | 0.0015042
| Rel.Std.Dev. |  | 0% | 43.3% |  |  | 18.17% | 5.906% | 29.78% | 7.906% | 2.599% | 5.386% | 3.114% | 2.67% | 0% | 0% | 12.91% | 7.27% | 0% | 4.919% | 2.242%
| PDM3 |  | 2.46% | -2.419% |  |  | -2.419% | -37.79% | 51.25% | 17.1% | 6.851% | 5.143% | -2.907% | -1.931% | 2.46% | 2.46% | 8.315% | -13.64% | -12.18% | 16.9% | -1.797%
| Z-Score (Absolute) |  | 0 | 4.07 | 3.37 | 156.78 | 0.56 | 4.64 | 5.62 | 1.69 | 0.51 | 0.31 | 0.62 | 0.51 | 0 | 0 | 0.67 | 1.85 | 1.69 | 1.66 | 0.49
| Analytical results for Ir in OREAS 238b (Indicative Value 0.001 ppm) | [27 empty] | Indicative
Ir | Round | Replicate | 1.01: Lab | [7 empty] | 1
| No | No | 1.01: 1.30 | [7 empty] | ppm
|  |  | 1.01: 4A*MS | [7 empty] | 3
3
| 1 | 1 | 1.01: 0.0009 | [7 empty] | 1
| 1 | 2 | 1.01: 0.0011 | [7 empty] | 3
| 1 | 3 | 1.01: < 0.0001 | [7 empty] | 16
| 1 | 4 | 1.01: 0.0004 | [7 empty] | 0.00065
| 1 | 5 | 1.01: 0.0008 | [7 empty] | 9
| Mean |  | 1.01: 0.0008
| Median |  | 1.01: 0.00085
| Std Dev. |  | 1.01: 0.00029439
| Rel.Std.Dev. |  | 1.01: 36.8%
| PDM3 |  | 1.01: 23.08%
Z-Score (Absolute) | bar: 1.01
| Analytical results for K in OREAS 238b (Certified Value 2.69 wt.%) | [27 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 2.7892 | 2.7654 | 2.69 | 3.11 | 2.525 | 2.75 | 2.59 | 2.71 | 2.7 | 2.66 | 2.67 | 2.69 | 2.65 | 2.8723 | 2.7016 | 2.7312 | 2.71 | 2.64 | 2.66 | 2.72 | 2.63 | 2.6475 | 3.0423 | 3.1523 |  |  | 1
| 1 | 2 | 2.7985 | 2.8038 | 2.73 | 3.21 | 2.565 | 2.75 | 2.6 | 2.78 | 2.7 | 2.7 | 2.72 | 2.64 | 2.64 | 2.549 | 2.705 | 2.7727 | 2.76 | 2.6 | 2.43 | 2.73 | 2.68 | 2.604 | 3.0404 | 3.1509 |  |  | 0
| 1 | 3 | 2.7817 | 2.7977 | 2.7 | 3.23 | 2.5745 | 2.761 | 2.59 | 2.65 | 2.75 | 2.72 | 2.74 | 2.69 | 2.58 | 2.7683 | 2.6474 | 3.138 | 2.64 | 2.58 | 2.59 | 2.74 | 2.69 | 2.6722 | 3.0435 | 3.1506 |  |  | 16
| 1 | 4 | 2.8775 | 2.8202 | 2.76 | 3.25 | 2.5596 | 2.758 | 2.62 | 2.79 | 2.74 | 2.77 | 2.68 | 2.68 | 2.62 | 2.7438 | 2.6644 | 2.7312 | 2.71 | 2.62 | 2.52 | 2.73 | 2.73 | 2.6519 | 3.1822 | 3.156 |  |  | 2.6916
| 1 | 5 | 2.8249 | 2.7742 | 2.74 | 3.25 | 2.5252 | 2.758 | 2.61 | 2.28 | 2.73 | 2.72 | 2.66 | 2.67 | 2.69 | 2.5694 | 2.6774 | 2.6648 | 2.64 | 2.62 | 2.66 | 2.73 | 2.68 | 2.6617 | 3.0422 | 3.153 |  |  | 37
| 1 | 6 | 2.8489 | 2.7502 | 2.72 | 3.37 | 2.6 | 2.753 | 2.58 | 2.76 | 2.74 | 2.85 | 2.83 | 2.73 | 2.66 | 2.2766 | 2.7073 | 2.7146 | 2.61 | 2.59 | 2.71 | 2.74 | 2.66 | 2.6146 | 2.9644 | 3.152
| Mean |  | 2.8201 | 2.7853 | 2.7233 | 3.2367 | 2.5582 | 2.755 | 2.5983 | 2.6617 | 2.7267 | 2.7367 | 2.7167 | 2.6833 | 2.64 | 2.6299 | 2.6839 | 2.7921 | 2.6783 | 2.6083 | 2.595 | 2.7317 | 2.6783 | 2.642 | 3.0525 | 3.1525
| Median |  | 2.8117 | 2.7859 | 2.725 | 3.24 | 2.5623 | 2.7555 | 2.595 | 2.735 | 2.735 | 2.72 | 2.7 | 2.685 | 2.645 | 2.6566 | 2.6895 | 2.7312 | 2.675 | 2.61 | 2.625 | 2.73 | 2.68 | 2.6497 | 3.0423 | 3.1521
| Std Dev. |  | 0.037532 | 0.026303 | 0.02582 | 0.083586 | 0.029163 | 0.0046476 | 0.01472 | 0.19405 | 0.021602 | 0.065929 | 0.063456 | 0.029439 | 0.037417 | 0.21256 | 0.024737 | 0.173 | 0.057067 | 0.022286 | 0.10445 | 0.0075277 | 0.033116 | 0.026919 | 0.070741 | 0.001947
| Rel.Std.Dev. |  | 1.331% | 0.9444% | 0.9481% | 2.582% | 1.14% | 0.1687% | 0.5665% | 7.291% | 0.7923% | 2.409% | 2.336% | 1.097% | 1.417% | 8.082% | 0.9217% | 6.196% | 2.131% | 0.8544% | 4.025% | 0.2756% | 1.236% | 1.019% | 2.317% | 0.06176%
| PDM3 |  | 4.777% | 3.481% | 1.181% | 20.25% | -4.954% | 2.357% | -3.463% | -1.11% | 1.305% | 1.676% | 0.9331% | -0.3053% | -1.915% | -2.291% | -0.2854% | 3.735% | -0.4911% | -3.092% | -3.587% | 1.49% | -0.4911% | -1.842% | 13.41% | 17.12%
| Z-Score (Absolute) |  | 1.36 | 0.97 | 0.26 | 6.1 | 1.62 | 0.62 | 1.16 | 0.44 | 0.3 | 0.41 | 0.19 | 0.19 | 0.69 | 0.8 | 0.19 | 1.04 | 0.25 | 1.05 | 1.2 | 0.36 | 0.25 | 0.66 | 4.01 | 5.14
| Analytical results for La in OREAS 238b (Certified Value 38.8 ppm) | [27 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
2
| 1 | 1 | 40.1 | 44.7819 | 41.3 | 43.4 | 35.69 | 35.8 | 39 | 41.7 | 41.7 | 40.4 | 40.1 | 43.6 | 36.4818 | 40 | 36.46 | 36.6 | 35.83 | 38.2 | 39.8 | 38.5 | 33.57 | 35.5012 |  |  |  |  | 1
| 1 | 2 | 40.2 | 45.1661 | 39.5 | 44.7 | 35.76 | 38.3 | 39 | 41 | 42.3 | 41.1 | 39.4 | 37.4 | 35.5466 | 41 | 36 | 37 | 35.99 | 39.7 | 36.5 | 39.3 | 33.09 | 34.0341 |  |  |  |  | 28
| 1 | 3 | 41.17 | 45.0644 | 40.8 | 44.3 | 35.1 | 39.9 | 39 | 39.2 | 41.8 | 39.3 | 40.6 | 39.1 | 36.8145 | 46 | 37.08 | 37.6 | 35.66 | 39.7 | 36.1 | 39.3 | 34 | 35.135 |  |  |  |  | 16
| 1 | 4 | 41.39 | 45.2 | 41.8 | 44.2 | 34.77 | 41.7 | 39 | 40.5 | 40.7 | 41 | 40.4 | 39.3 | 35.2018 | 40 | 35.34 | 35.6 | 36.26 | 39.5 | 34.3 | 40.1 | 33.57 | 35.1388 |  |  |  |  | 38.8276
| 1 | 5 | 41.17 | 44.8723 | 39 | 44.1 | 35.66 | 40.4 | 39 | 40.6 | 41.6 | 39.2 | 39.7 | 38 | 35.8355 | 39 | 36.35 | 36.2 | 36.37 | 38.7 | 38.3 | 37.7 | 33.86 | 34.4192 |  |  |  |  | 38
| 1 | 6 | 41.02 | 44.8271 | 40.2 | 43.9 | 35.37 | 40 | 39 | 40.7 | 38.5 | 41.4 | 40.8 | 36.8 | 36.6768 | 40 | 36.93 | 36.1 | 36.24 | 37.8 | 38.6 | 38.1 | 33.21 | 34.9386
| Mean |  | 40.8417 | 44.9853 | 40.4333 | 44.1 | 35.3917 | 39.35 | 39 | 40.6167 | 41.1 | 40.4 | 40.1667 | 39.0333 | 36.0929 | 41 | 36.36 | 36.5167 | 36.0583 | 38.9333 | 37.2667 | 38.8333 | 33.55 | 34.8612
| Median |  | 41.095 | 44.9684 | 40.5 | 44.15 | 35.515 | 39.95 | 39 | 40.65 | 41.65 | 40.7 | 40.25 | 38.55 | 36.1587 | 40 | 36.405 | 36.4 | 36.115 | 39.1 | 37.4 | 38.9 | 33.57 | 35.0368
| Std Dev. |  | 0.54952 | 0.18122 | 1.0708 | 0.43359 | 0.39209 | 2.0521 | 0 | 0.81833 | 1.3755 | 0.94868 | 0.53914 | 2.4353 | 0.65902 | 2.5298 | 0.63614 | 0.7111 | 0.27795 | 0.82138 | 2.0007 | 0.89144 | 0.35401 | 0.53798
| Rel.Std.Dev. |  | 1.345% | 0.4028% | 2.648% | 0.9832% | 1.108% | 5.215% | 0% | 2.015% | 3.347% | 2.348% | 1.342% | 6.239% | 1.826% | 6.17% | 1.75% | 1.947% | 0.7708% | 2.11% | 5.369% | 2.296% | 1.055% | 1.543%
| PDM3 |  | 5.187% | 15.86% | 4.135% | 13.58% | -8.849% | 1.345% | 0.444% | 4.608% | 5.852% | 4.05% | 3.449% | 0.5298% | -7.043% | 5.595% | -6.355% | -5.952% | -7.132% | 0.2723% | -4.02% | 0.01471% | -13.59% | -10.22%
| Z-Score (Absolute) |  | 0.65 | 2.11 | 0.5 | 1.8 | 1.28 | 0.12 | 0.01 | 0.57 | 0.74 | 0.49 | 0.41 | 0.01 | 1.04 | 0.7 | 0.94 | 0.89 | 1.05 | 0.03 | 0.62 | 0.06 | 1.94 | 1.47
| Analytical results for Li in OREAS 238b (Certified Value 48.8 ppm) | [27 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 49.7 | 53.28 | 50 | 53 | 48.77 | 47.7 | 51 | 47.8 | 76 | 51.9 | 45.5 | 51.4 | 43.8 | 52 | 50.4272 | 50 | 41.24 | 48.9 | 51.66 | 46.9 | 45.7 | 49.1 | 52.6 | 41.0236 |  |  | 1
| 1 | 2 | 48.5 | 54.22 | 50 | 53 | 46.8733 | 47.3 | 51 | 48.2 | 67.9 | 53.3 | 47.1 | 52.2 | 44.4 | 51 | 49.1345 | 51 | 41.07 | 48.1 | 51.03 | 47.3 | 43.5 | 49 | 51.6 | 42.5055 |  |  | 0
| 1 | 3 | 49.8 | 53.27 | 50 | 53 | 48.515 | 46.1 | 51 | 47.2 | 70.4 | 52.5 | 47.2 | 50.8 | 45 | 50.8 | 50.6463 | 57 | 42.31 | 44.8 | 50.53 | 47.5 | 42.7 | 50.4 | 53.2 | 41.4519 |  |  | 16
| 1 | 4 | 49.5 | 53.83 | 50 | 53 | 48.61 | 45.7 | 51 | 50.2 | 73.2 | 54.2 | 45.8 | 51 | 43.6 | 51.5 | 52.3665 | 50 | 41.92 | 48.6 | 51.25 | 47.7 | 40.5 | 50.7 | 52.5 | 41.729 |  |  | 48.7761
| 1 | 5 | 48.9 | 53.22 | 50 | 52 | 49.82 | 45.6 | 51 | 45.7 | 68.9 | 53.2 | 44.4 | 49.6 | 46.8 | 50.9 | 50.3353 | 49 | 43.76 | 50.2 | 51.92 | 47.2 | 44.7 | 49.9 | 52.8 | 36.803 |  |  | 39
| 1 | 6 | 50.3 | 53.55 | 50 | 53 | 47.5033 | 46.6 | 50 | 46.8 | 68.2 | 54.7 | 45.1 | 52.7 | 47 | 50.7 | 51.6081 | 50 | 41.76 | 48 | 49.59 | 45.2 | 44.9 | 51 | 51.7 | 39.7485
| Mean |  | 49.45 | 53.5617 | 50 | 52.8333 | 48.3486 | 46.5 | 50.8333 | 47.65 | 70.7667 | 53.3 | 45.85 | 51.2833 | 45.1 | 51.15 | 50.753 | 51.1667 | 42.01 | 48.1 | 50.9967 | 46.9667 | 43.6667 | 50.0167 | 52.4 | 40.5436
| Median |  | 49.6 | 53.415 | 50 | 53 | 48.5625 | 46.35 | 51 | 47.5 | 69.65 | 53.25 | 45.65 | 51.2 | 44.7 | 50.95 | 50.5368 | 50 | 41.84 | 48.35 | 51.14 | 47.25 | 44.1 | 50.15 | 52.55 | 41.2377
| Std Dev. |  | 0.65038 | 0.39666 | 0 | 0.40825 | 1.0324 | 0.86023 | 0.40825 | 1.5202 | 3.2166 | 1.0373 | 1.1113 | 1.0962 | 1.4792 | 0.501 | 1.1177 | 2.9269 | 0.96949 | 1.8 | 0.84313 | 0.90701 | 1.8822 | 0.83287 | 0.62929 | 2.0462
| Rel.Std.Dev. |  | 1.315% | 0.7406% | 0% | 0.7727% | 2.135% | 1.85% | 0.8031% | 3.19% | 4.545% | 1.946% | 2.424% | 2.138% | 3.28% | 0.9795% | 2.202% | 5.72% | 2.308% | 3.742% | 1.653% | 1.931% | 4.31% | 1.665% | 1.201% | 5.047%
| PDM3 |  | 1.382% | 9.811% | 2.509% | 8.318% | -0.8765% | -4.666% | 4.218% | -2.309% | 45.08% | 9.275% | -5.999% | 5.14% | -7.537% | 4.867% | 4.053% | 4.901% | -13.87% | -1.386% | 4.552% | -3.71% | -10.48% | 2.543% | 7.43% | -16.88%
| Z-Score (Absolute) |  | 0.16 | 1 | NA | 0.8 | 0.47 | 0.99 | 0.23 | 0.67 | 5.87 | 0.93 | 1.18 | 0.36 | 1.39 | 0.32 | 0.21 | 0.33 | 2.27 | 0.54 | 0.28 | 0.86 | 1.8 | 0 | 0.67 | 2.68
Results from laboratory 1.05 were removed due to their 10 ppm reading resolution.
| Analytical results for Lu in OREAS 238b (Certified Value 0.29 ppm) | [27 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.20; 1.08: 1.21; 1.09: 1.26; 1.10: 1.28; 1.11: 1.30; 1.12: 1.34 |  |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*OES/MS; 1.11: 4A*MS; 1.12: 4A*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.25441; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.226; 1.08: 0.25; 1.09: 0.31; 1.10: 0.2; 1.11: 0.1249; 1.12: 0.38472 |  |  | 1
| 1 | 2 | 1.01: 0.25497; 1.03: 0.3; 1.05: 0.4; 1.06: 0.2; 1.07: 0.224; 1.08: 0.27; 1.09: 0.33; 1.10: 0.2; 1.11: 0.1328; 1.12: 0.35774 |  |  | 29
| 1 | 3 | 1.01: 0.26578; 1.03: 0.28; 1.05: 0.4; 1.06: 0.3; 1.07: 0.233; 1.08: 0.27; 1.09: 0.34; 1.10: 0.2; 1.11: 0.1093; 1.12: 0.37326 |  |  | 16
| 1 | 4 | 1.01: 0.25153; 1.03: 0.3; 1.05: 0.3; 1.06: 0.2; 1.07: 0.229; 1.08: 0.26; 1.09: 0.35; 1.10: 0.2; 1.11: 0.1017; 1.12: 0.3622 |  |  | 0.2916
| 1 | 5 | 1.01: 0.25654; 1.03: 0.32; 1.05: 0.3; 1.06: 0.3; 1.07: 0.221; 1.08: 0.28; 1.09: 0.33; 1.10: 0.2; 1.11: 0.1214; 1.12: 0.35815 |  |  | 40
| 1 | 6 | 1.01: 0.26834; 1.03: 0.32; 1.05: 0.4; 1.06: 0.3; 1.07: 0.228; 1.08: 0.27; 1.09: 0.33; 1.10: 0.2; 1.12: 0.36103
| Mean |  | 1.01: 0.25859; 1.03: 0.30333; 1.05: 0.35; 1.06: 0.26667; 1.07: 0.22683; 1.08: 0.26667; 1.09: 0.33167; 1.10: 0.2; 1.11: 0.11802; 1.12: 0.36618
| Median |  | 1.01: 0.25576; 1.03: 0.3; 1.05: 0.35; 1.06: 0.3; 1.07: 0.227; 1.08: 0.27; 1.09: 0.33; 1.10: 0.2; 1.11: 0.1214; 1.12: 0.36162
| Std Dev. |  | 1.01: 0.006803; 1.03: 0.015055; 1.05: 0.054772; 1.06: 0.05164; 1.07: 0.0041673; 1.08: 0.010328; 1.09: 0.013292; 1.10: 3.0405e-17; 1.11: 0.012446; 1.12: 0.010695
| Rel.Std.Dev. |  | 1.01: 2.631%; 1.03: 4.963%; 1.05: 15.65%; 1.06: 19.36%; 1.07: 1.837%; 1.08: 3.873%; 1.09: 4.008%; 1.10: 1.52e-14%; 1.11: 10.55%; 1.12: 2.921%
| PDM3 |  | 1.01: -11.32%; 1.03: 4.026%; 1.05: 20.03%; 1.06: -8.549%; 1.07: -22.21%; 1.08: -8.549%; 1.09: 13.74%; 1.10: -31.41%; 1.11: -59.53%; 1.12: 25.58%
| Z-Score (Absolute) |  | 1.01: 0.14; 1.03: 0.62; 1.05: NA; 1.06: NA; 1.07: 0.67; 1.08: 0; 1.09: 1.1; 1.10: NA; 1.11: 2.52; 1.12: 1.68
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 238b (Certified Value 1.74 wt.%) | [27 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 2
3
| 1 | 1 | 1.6458 | 1.7685 | 1.76 | 1.85 | 1.74 | 1.785 | 1.7 | 1.76 | 1.7 | 1.7 | 1.68 | 1.72 | 1.67 | 1.7816 | 1.743 | 1.8213 | 1.73 | 1.76 | 1.82 | 1.71 | 1.7 | 1.7318 | 2.0753 | 1.6579 |  |  | 1
| 1 | 2 | 1.6603 | 1.8187 | 1.74 | 1.84 | 1.76 | 1.781 | 1.75 | 1.8 | 1.7 | 1.72 | 1.74 | 1.68 | 1.64 | 1.7121 | 1.7833 | 1.8514 | 1.71 | 1.76 | 1.83 | 1.73 | 1.7 | 1.7132 | 2.0973 | 1.6844 |  |  | 0
| 1 | 3 | 1.6879 | 1.8128 | 1.75 | 1.88 | 1.74 | 1.788 | 1.75 | 1.77 | 1.7 | 1.71 | 1.74 | 1.69 | 1.62 | 1.7336 | 1.7289 | 2.1107 | 1.73 | 1.77 | 1.82 | 1.72 | 1.72 | 1.763 | 2.0703 | 1.7072 |  |  | 16
| 1 | 4 | 1.7007 | 1.8332 | 1.79 | 1.83 | 1.75 | 1.784 | 1.78 | 1.8 | 1.71 | 1.76 | 1.69 | 1.69 | 1.67 | 1.7434 | 1.7615 | 1.8152 | 1.76 | 1.75 | 1.83 | 1.72 | 1.71 | 1.7444 | 2.1644 | 1.6073 |  |  | 1.7445
| 1 | 5 | 1.7053 | 1.816 | 1.8 | 1.85 | 1.78 | 1.787 | 1.77 | 1.73 | 1.72 | 1.72 | 1.68 | 1.69 | 1.7 | 1.7395 | 1.7568 | 1.767 | 1.75 | 1.76 | 1.77 | 1.71 | 1.7 | 1.7487 | 2.0633 | 1.6356 |  |  | 41
| 1 | 6 | 1.7067 | 1.783 | 1.81 | 1.84 | 1.78 | 1.782 | 1.74 | 1.78 | 1.73 | 1.77 | 1.75 | 1.72 | 1.69 | 1.715 | 1.7802 | 1.8092 | 1.7 | 1.77 | 1.82 | 1.72 | 1.7 | 1.7157 | 2.0007 | 1.6334
| Mean |  | 1.6845 | 1.8054 | 1.775 | 1.8483 | 1.7583 | 1.7845 | 1.7483 | 1.7733 | 1.71 | 1.73 | 1.7133 | 1.6983 | 1.665 | 1.7375 | 1.759 | 1.8625 | 1.73 | 1.7617 | 1.815 | 1.7183 | 1.705 | 1.7361 | 2.0785 | 1.6543
| Median |  | 1.6943 | 1.8144 | 1.775 | 1.845 | 1.755 | 1.7845 | 1.75 | 1.775 | 1.705 | 1.72 | 1.715 | 1.69 | 1.67 | 1.7366 | 1.7592 | 1.8182 | 1.73 | 1.76 | 1.82 | 1.72 | 1.7 | 1.7381 | 2.0728 | 1.6467
| Std Dev. |  | 0.025623 | 0.024407 | 0.02881 | 0.017224 | 0.018348 | 0.0027386 | 0.027869 | 0.026583 | 0.012649 | 0.028284 | 0.033267 | 0.017224 | 0.030166 | 0.025101 | 0.021028 | 0.1246 | 0.022804 | 0.0075277 | 0.022583 | 0.0075277 | 0.0083666 | 0.019546 | 0.053092 | 0.036641
| Rel.Std.Dev. |  | 1.521% | 1.352% | 1.623% | 0.9319% | 1.044% | 0.1535% | 1.594% | 1.499% | 0.7397% | 1.635% | 1.942% | 1.014% | 1.812% | 1.445% | 1.195% | 6.69% | 1.318% | 0.4273% | 1.244% | 0.4381% | 0.4907% | 1.126% | 2.554% | 2.215%
| PDM3 |  | -3.444% | 3.488% | 1.747% | 5.951% | 0.7915% | 2.291% | 0.2183% | 1.651% | -1.979% | -0.8326% | -1.788% | -2.648% | -4.559% | -0.4008% | 0.8277% | 6.761% | -0.8326% | 0.9826% | 4.04% | -1.501% | -2.266% | -0.481% | 19.15% | -5.172%
| Z-Score (Absolute) |  | 1.21 | 1.3 | 0.67 | 2.19 | 0.32 | 0.86 | 0.11 | 0.63 | 0.68 | 0.27 | 0.61 | 0.93 | 1.62 | 0.11 | 0.33 | 2.48 | 0.27 | 0.39 | 1.5 | 0.51 | 0.79 | 0.14 | 6.96 | 1.84
| Analytical results for Mn in OREAS 238b (Certified Value 0.039 wt.%) | [27 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0404 | 0.039291 | 0.0396 | 0.0419 | 0.042305 | 0.0394 | 0.0397 | 0.0361 | 0.0388 | 0.0377 | 0.0379 | 0.0383 | 0.0388 | 0.04 | 0.039333 | 0.0372 | 0.0366 | 0.03706 | 0.0392 | 0.0373 | 0.0382 | 0.040359 | 0.0422 | 0.0385 |  |  | 1
| 1 | 2 | 0.0397 | 0.039747 | 0.04 | 0.0432 | 0.042413 | 0.04 | 0.038 | 0.0368 | 0.0394 | 0.038 | 0.0387 | 0.0376 | 0.0378 | 0.0386 | 0.039463 | 0.0379 | 0.0369 | 0.03687 | 0.0396 | 0.0376 | 0.0384 | 0.039919 | 0.0419 | 0.039 |  |  | 30
| 1 | 3 | 0.0413 | 0.039525 | 0.0402 | 0.0411 | 0.041604 | 0.039 | 0.0386 | 0.0363 | 0.0398 | 0.038 | 0.0392 | 0.0377 | 0.0373 | 0.0387 | 0.039323 | 0.0426 | 0.0361 | 0.0367 | 0.0396 | 0.0377 | 0.0388 | 0.040733 | 0.0421 | 0.0396 |  |  | 16
| 1 | 4 | 0.0421 | 0.039506 | 0.0406 | 0.0407 | 0.041581 | 0.0383 | 0.0395 | 0.0368 | 0.0399 | 0.0393 | 0.0378 | 0.0378 | 0.0387 | 0.0393 | 0.039037 | 0.0372 | 0.0359 | 0.03699 | 0.0395 | 0.0376 | 0.0384 | 0.040172 | 0.044 | 0.0377 |  |  | 0.039061
| 1 | 5 | 0.0421 | 0.039546 | 0.0406 | 0.0417 | 0.041974 | 0.0385 | 0.039 | 0.0354 | 0.0402 | 0.0385 | 0.0374 | 0.0382 | 0.039 | 0.0389 | 0.039463 | 0.0364 | 0.0364 | 0.0368 | 0.0381 | 0.0374 | 0.0385 | 0.040392 | 0.0415 | 0.0383 |  |  | 42
| 1 | 6 | 0.0414 | 0.039106 | 0.0416 | 0.0434 | 0.042159 | 0.0401 | 0.0377 | 0.0363 | 0.0407 | 0.0401 | 0.039 | 0.0386 | 0.0385 | 0.0383 | 0.03962 | 0.0372 | 0.0362 | 0.03682 | 0.0393 | 0.0375 | 0.0384 | 0.040227 | 0.0406 | 0.0382
| Mean |  | 0.041167 | 0.039454 | 0.040433 | 0.042 | 0.042006 | 0.039217 | 0.03875 | 0.036283 | 0.0398 | 0.0386 | 0.038333 | 0.038033 | 0.03835 | 0.038967 | 0.039373 | 0.038083 | 0.03635 | 0.036873 | 0.039217 | 0.037517 | 0.03845 | 0.0403 | 0.04205 | 0.03855
| Median |  | 0.04135 | 0.039516 | 0.0404 | 0.0418 | 0.042066 | 0.0392 | 0.0388 | 0.0363 | 0.03985 | 0.03825 | 0.0383 | 0.038 | 0.0386 | 0.0388 | 0.039398 | 0.0372 | 0.0363 | 0.036845 | 0.0394 | 0.03755 | 0.0384 | 0.040293 | 0.042 | 0.0384
| Std Dev. |  | 0.00095429 | 0.00022349 | 0.00068605 | 0.0010954 | 0.00035258 | 0.00075211 | 0.00080187 | 0.00051929 | 0.00065422 | 0.0009252 | 0.00073121 | 0.00039328 | 0.00065955 | 0.00060553 | 0.00019695 | 0.0022631 | 0.00036194 | 0.00013171 | 0.00057067 | 0.0001472 | 0.00019748 | 0.00027067 | 0.0011185 | 0.00066558
| Rel.Std.Dev. |  | 2.318% | 0.5665% | 1.697% | 2.608% | 0.8394% | 1.918% | 2.069% | 1.431% | 1.644% | 2.397% | 1.908% | 1.034% | 1.72% | 1.554% | 0.5002% | 5.943% | 0.9957% | 0.3572% | 1.455% | 0.3923% | 0.5136% | 0.6716% | 2.66% | 1.727%
| PDM3 |  | 5.39% | 1.004% | 3.513% | 7.523% | 7.539% | 0.3978% | -0.7969% | -7.112% | 1.891% | -1.181% | -1.864% | -2.632% | -1.821% | -0.2422% | 0.7988% | -2.504% | -6.941% | -5.601% | 0.3978% | -3.954% | -1.565% | 3.172% | 7.651% | -1.309%
| Z-Score (Absolute) |  | 1.94 | 0.5 | 1.32 | 2.64 | 2.64 | 0.3 | 0.09 | 2.16 | 0.79 | 0.22 | 0.44 | 0.69 | 0.43 | 0.09 | 0.43 | 0.66 | 2.11 | 1.67 | 0.3 | 1.13 | 0.34 | 1.21 | 2.68 | 0.26
| Analytical results for Mo in OREAS 238b (Certified Value 0.9 ppm) | [27 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.9 | 1 | < 1 | < 2 | 0.9 | 0.83 | 1.3 | 2.9 | 0.89 | 0.85 | 0.96 | 0.86 | 0.9 | 0.96801 | 1.1 | 1.344 | 0.94 | < 1 | 0.94 | 1 | 0.84 | 0.7334 | < 5 | 2.5133 | 0.87247 |  | 1
| 1 | 2 | 0.8 | 1 | < 1 | < 2 | 0.9 | 0.87 | 0.9 | 2.8 | 0.92 | 0.89 | 0.91 | 0.87 | 0.85 | 0.94405 | 0.97 | 1.447 | 0.93 | < 1 | 0.93 | 0.87 | 0.82 | 0.7345 | < 5 | 2.9583 | 0.91315 |  | 31
| 1 | 3 | 1 | 1 | < 1 | < 2 | 0.9 | 0.86 | 0.9 | 3.1 | 0.86 | 0.87 | 0.92 | 0.86 | 0.88 | 0.95433 | 0.94 | 1.754 | 0.87 | < 1 | 0.86 | 0.92 | 0.86 | 0.7226 | < 5 | 2.7914 | 0.89795 |  | 16
| 1 | 4 | 0.9 | 1 | < 1 | < 2 | 0.9 | 0.82 | 0.9 | 3 | 0.86 | 0.95 | 0.95 | 0.92 | 0.95 | 0.95048 | 0.96 | 1.633 | 0.97 | < 1 | 1.02 | 0.87 | 0.85 | 0.7539 | < 5 | 2.2352 | 0.92711 |  | 0.90149
| 1 | 5 | 0.9 | 1 | < 1 | < 2 | 0.9 | 0.83 | 0.7 | 2.9 | 0.82 | 0.87 | 0.89 | 0.93 | 0.91 | 0.96902 | 0.9 | 1.567 | 0.93 | < 1 | 0.92 | 0.95 | 0.85 | 0.6686 | < 5 | 2.5412 | 0.85347 |  | 43
| 1 | 6 | 1 | 1 | < 1 | < 2 | 0.9 | 0.89 | 1.1 | 2.8 | 0.84 | 0.87 | 0.97 | 0.89 | 0.9 | 0.94698 | 0.94 | 1.285 | 0.9 | < 1 | 0.77 | 0.96 | 0.85 | 0.7366 | < 5 | 2.8749 | 0.91377
| Mean |  | 0.91667 | 1 |  |  | 0.9 | 0.85 | 0.96667 | 2.9167 | 0.865 | 0.88333 | 0.93333 | 0.88833 | 0.89833 | 0.95548 | 0.96833 | 1.505 | 0.92333 |  | 0.90667 | 0.92833 | 0.845 | 0.72493 |  | 2.6524 | 0.89632
| Median |  | 0.9 | 1 |  |  | 0.9 | 0.845 | 0.9 | 2.9 | 0.86 | 0.87 | 0.935 | 0.88 | 0.9 | 0.9524 | 0.95 | 1.507 | 0.93 |  | 0.925 | 0.935 | 0.85 | 0.73395 |  | 2.6663 | 0.90555
| Std Dev. |  | 0.075277 | 0 |  |  | 0 | 0.027568 | 0.20656 | 0.1169 | 0.035637 | 0.035024 | 0.031411 | 0.030605 | 0.033116 | 0.010673 | 0.068823 | 0.17884 | 0.034448 |  | 0.084301 | 0.051929 | 0.013784 | 0.029385 |  | 0.27136 | 0.028083
| Rel.Std.Dev. |  | 8.212% | 0% |  |  | 0% | 3.243% | 21.37% | 4.008% | 4.12% | 3.965% | 3.365% | 3.445% | 3.686% | 1.117% | 7.107% | 11.88% | 3.731% |  | 9.298% | 5.594% | 1.631% | 4.053% |  | 10.23% | 3.133%
| PDM3 |  | 1.684% | 10.93% |  |  | -0.165% | -5.711% | 7.23% | 223.5% | -4.047% | -2.014% | 3.533% | -1.459% | -0.3499% | 5.989% | 7.415% | 66.95% | 2.423% |  | 0.5745% | 2.978% | -6.266% | -19.58% |  | 194.2% | -0.5732%
| Z-Score (Absolute) |  | 0.07 | NA | 5.46 | 1.17 | 0.15 | 0.82 | 0.73 | 26.6 | 0.62 | 0.38 | 0.29 | 0.31 | 0.18 | 0.58 | 0.75 | 7.87 | 0.15 | 5.46 | 0.07 | 0.22 | 0.88 | 2.48 | 21.07 | 23.09 | 0.2
Results from laboratory 1.05 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Na in OREAS 238b (Certified Value 0.676 wt.%) | [27 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.6861 | 0.7141 | 0.7 | 0.76 | 0.735 | 0.664 | 0.66 | 0.688 | 0.67 | 0.67 | 0.66 | 0.68 | 0.69 | 0.692 | 0.64839 | 0.675 | 0.67 | 0.71 | 0.67 | 0.66 | 0.67 | 0.67169 | 0.7134 |  |  |  | 1
| 1 | 2 | 0.6801 | 0.72256 | 0.69 | 0.76 | 0.73039 | 0.686 | 0.66 | 0.704 | 0.68 | 0.66 | 0.68 | 0.66 | 0.68 | 0.664 | 0.64062 | 0.682 | 0.69 | 0.7 | 0.68 | 0.66 | 0.667 | 0.65983 | 0.7062 |  |  |  | 0
| 1 | 3 | 0.6846 | 0.72096 | 0.69 | 0.78 | 0.72033 | 0.655 | 0.66 | 0.691 | 0.66 | 0.67 | 0.68 | 0.67 | 0.67 | 0.669 | 0.6523 | 0.772 | 0.65 | 0.7 | 0.68 | 0.66 | 0.67 | 0.68159 | 0.7266 |  |  |  | 16
| 1 | 4 | 0.696 | 0.72394 | 0.7 | 0.74 | 0.7205 | 0.672 | 0.7 | 0.706 | 0.65 | 0.69 | 0.66 | 0.67 | 0.68 | 0.672 | 0.65563 | 0.668 | 0.68 | 0.7 | 0.68 | 0.66 | 0.674 | 0.67218 | 0.6916 |  |  |  | 0.67587
| 1 | 5 | 0.693 | 0.71621 | 0.7 | 0.78 | 0.7205 | 0.653 | 0.69 | 0.679 | 0.66 | 0.67 | 0.66 | 0.66 | 0.7 | 0.672 | 0.66204 | 0.653 | 0.66 | 0.7 | 0.66 | 0.66 | 0.662 | 0.67375 | 0.6943 |  |  |  | 44
| 1 | 6 | 0.6916 | 0.71973 | 0.7 | 0.76 | 0.726 | 0.685 | 0.66 | 0.697 | 0.65 | 0.69 | 0.69 | 0.68 | 0.69 | 0.649 | 0.65149 | 0.668 | 0.65 | 0.71 | 0.68 | 0.66 | 0.668 | 0.6614 | 0.6997
| Mean |  | 0.68857 | 0.71958 | 0.69667 | 0.76333 | 0.72545 | 0.66917 | 0.67167 | 0.69417 | 0.66167 | 0.675 | 0.67167 | 0.67 | 0.685 | 0.66967 | 0.65174 | 0.68633 | 0.66667 | 0.70333 | 0.675 | 0.66 | 0.6685 | 0.67007 | 0.7053
| Median |  | 0.68885 | 0.72035 | 0.7 | 0.76 | 0.72325 | 0.668 | 0.66 | 0.694 | 0.66 | 0.67 | 0.67 | 0.67 | 0.685 | 0.6705 | 0.6519 | 0.6715 | 0.665 | 0.7 | 0.68 | 0.66 | 0.669 | 0.67194 | 0.70295
| Std Dev. |  | 0.0059601 | 0.0037737 | 0.005164 | 0.015055 | 0.006182 | 0.014359 | 0.018348 | 0.010226 | 0.01169 | 0.012247 | 0.013292 | 0.0089443 | 0.010488 | 0.013924 | 0.0071618 | 0.043057 | 0.01633 | 0.005164 | 0.0083666 | 0 | 0.0039875 | 0.008164 | 0.013116
| Rel.Std.Dev. |  | 0.8656% | 0.5244% | 0.7412% | 1.972% | 0.8522% | 2.146% | 2.732% | 1.473% | 1.767% | 1.814% | 1.979% | 1.335% | 1.531% | 2.079% | 1.099% | 6.273% | 2.449% | 0.7342% | 1.239% | 0% | 0.5965% | 1.218% | 1.86%
| PDM3 |  | 1.879% | 6.468% | 3.078% | 12.94% | 7.337% | -0.9913% | -0.6214% | 2.708% | -2.101% | -0.1282% | -0.6214% | -0.868% | 1.351% | -0.9173% | -3.569% | 1.549% | -1.361% | 4.064% | -0.1282% | -2.348% | -1.09% | -0.8569% | 4.355%
| Z-Score (Absolute) |  | 0.82 | 2.68 | 1.3 | 5.32 | 3.04 | 0.35 | 0.2 | 1.15 | 0.8 | 0 | 0.2 | 0.3 | 0.6 | 0.32 | 1.4 | 0.67 | 0.5 | 1.71 | 0 | 0.9 | 0.39 | 0.3 | 1.82
| Analytical results for Nb in OREAS 238b (Certified Value 14.3 ppm) | [27 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 15.79 | 12.9809 | 1 | 14.62 | 14 | 10 | 12.5 | 13.4 | 15.1 | 14.9 | 14.8 | 13.9 | 14.35 | 14.831 | 15.8 | 12.7 | 15.45 | 14.8 | 11.81 | 10.1037 | 15.2975 | [5 empty] | 1
| 1 | 2 | 15.45 | 12.952 | 1.8 | 15.76 | 13.9 | 9 | 13.8 | 13.8 | 15.1 | 14.9 | 14.6 | 14 | 14.42 | 14.9498 | 14.8 | 13.7 | 15.08 | 13.6 | 12.22 | 10.2466 | 15.7651 | [5 empty] | 32
| 1 | 3 | 16.06 | 12.9614 | 0.2 | 14.65 | 15.7 | 10 | 13.1 | 13.9 | 14.2 | 15.7 | 14 | 14.2 | 14.56 | 14.9513 | 15.1 | 12.7 | 15.3 | 13.5 | 11.95 | 10.3268 | 15.3818 | [5 empty] | 16
| 1 | 4 | 15.53 | 13.0932 | 0.5 | 14.97 | 14.4 | 9 | 13.9 | 14.1 | 15.1 | 14.8 | 15 | 14.1 | 14.41 | 15.1842 | 14.8 | 12.3 | 15.42 | 13.1 | 12.35 | 10.2904 | 15.4103 | [5 empty] | 14.3051
| 1 | 5 | 14.64 | 13.0259 | 0.9 | 16.06 | 15.4 | 10 | 11.6 | 14.5 | 14.5 | 14.5 | 14.2 | 14.3 | 15.03 | 14.8993 | 14.8 | 13.4 | 15.22 | 14.4 | 12.43 | 10.1504 | 15.2324 | [5 empty] | 45
| 1 | 6 | 15.45 | 12.9548 | 2 | 15.26 | 13.6 | 10 | 12.6 | 14.3 | 14.8 | 15 | 14.8 | 14.3 | 14.83 | 14.4958 | 14.8 | 12.9 | 14.73 | 14.1 | 12.32 | 9.9183 | 15.6832
| Mean |  | 15.4867 | 12.9947 | 1.0667 | 15.22 | 14.5 | 9.6667 | 12.9167 | 14 | 14.8 | 14.9667 | 14.5667 | 14.1333 | 14.6 | 14.8853 | 15.0167 | 12.95 | 15.2 | 13.9167 | 12.18 | 10.1727 | 15.4617
| Median |  | 15.49 | 12.9711 | 0.95 | 15.115 | 14.2 | 10 | 12.85 | 14 | 14.95 | 14.9 | 14.7 | 14.15 | 14.49 | 14.9246 | 14.8 | 12.8 | 15.26 | 13.85 | 12.27 | 10.1985 | 15.3961
| Std Dev. |  | 0.47794 | 0.055474 | 0.70899 | 0.59097 | 0.8579 | 0.5164 | 0.87044 | 0.38987 | 0.37947 | 0.39833 | 0.38816 | 0.1633 | 0.27218 | 0.22482 | 0.40208 | 0.51284 | 0.26706 | 0.63061 | 0.24593 | 0.15036 | 0.21432
| Rel.Std.Dev. |  | 3.086% | 0.4269% | 66.47% | 3.883% | 5.917% | 5.342% | 6.739% | 2.785% | 2.564% | 2.661% | 2.665% | 1.155% | 1.864% | 1.51% | 2.678% | 3.96% | 1.757% | 4.531% | 2.019% | 1.478% | 1.386%
| PDM3 |  | 8.26% | -9.16% | -92.54% | 6.396% | 1.363% | -32.42% | -9.706% | -2.133% | 3.46% | 4.625% | 1.829% | -1.201% | 2.062% | 4.056% | 4.974% | -9.473% | 6.256% | -2.715% | -14.86% | -28.89% | 8.086%
| Z-Score (Absolute) |  | 1 | 1.62 | 14.15 | 0.72 | 0.04 | NA | 1.7 | 0.56 | 0.28 | 0.46 | 0.04 | 0.42 | 0.07 | 0.37 | 0.51 | 1.66 | 0.7 | 0.65 | 2.47 | 4.58 | 0.98
Results from laboratory 1.09 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 238b (Certified Value 34.1 ppm) | [27 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.03; 1.05: 1.05; 1.06: 1.06; 1.07: 1.16; 1.08: 1.19; 1.09: 1.20; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS |  | 1
2
| 1 | 1 | 1.01: 33.2431; 1.05: 35; 1.06: 35; 1.07: 34.6; 1.08: 36.4; 1.09: 33.11; 1.10: 34.21; 1.11: 33.5; 1.12: 29.87; 1.13: 32.8415 |  | 1
| 1 | 2 | 1.01: 32.8328; 1.05: 35.2; 1.06: 37; 1.07: 29.3; 1.08: 35.2; 1.09: 32.98; 1.10: 34.75; 1.11: 34.2; 1.12: 29.27; 1.13: 30.7397 |  | 9
| 1 | 3 | 1.01: 33.5895; 1.05: 34.2; 1.06: 36.5; 1.07: 39.3; 1.08: 35.3; 1.09: 34.11; 1.10: 35.41; 1.11: 34.4; 1.12: 29.92; 1.13: 31.3544 |  | 16
| 1 | 4 | 1.01: 34.0179; 1.05: 35.5; 1.06: 35.9; 1.07: 31; 1.08: 34.4; 1.09: 33.22; 1.10: 34.6; 1.11: 34.5; 1.12: 29.77; 1.13: 31.4163 |  | 34.1112
| 1 | 5 | 1.01: 33.8499; 1.05: 35.2; 1.06: 36.6; 1.07: 34.9; 1.08: 34; 1.09: 33.77; 1.10: 33.86; 1.11: 33.3; 1.12: 29.77; 1.13: 31.0144 |  | 46
| 1 | 6 | 1.01: 34.0362; 1.05: 36.6; 1.06: 36.2; 1.07: 34.5; 1.08: 34.2; 1.09: 34.22; 1.10: 34.06; 1.11: 33; 1.12: 29.11; 1.13: 31.5037
| Mean |  | 1.01: 33.5949; 1.05: 35.2833; 1.06: 36.2; 1.07: 33.9333; 1.08: 34.9167; 1.09: 33.5683; 1.10: 34.4817; 1.11: 33.8167; 1.12: 29.6183; 1.13: 31.4783
| Median |  | 1.01: 33.7197; 1.05: 35.2; 1.06: 36.35; 1.07: 34.55; 1.08: 34.8; 1.09: 33.495; 1.10: 34.405; 1.11: 33.85; 1.12: 29.77; 1.13: 31.3854
| Std Dev. |  | 1.01: 0.47817; 1.05: 0.78081; 1.06: 0.6957; 1.07: 3.4806; 1.08: 0.89981; 1.09: 0.53596; 1.10: 0.56318; 1.11: 0.63061; 1.12: 0.34061; 1.13: 0.7265
| Rel.Std.Dev. |  | 1.01: 1.423%; 1.05: 2.213%; 1.06: 1.922%; 1.07: 10.26%; 1.08: 2.577%; 1.09: 1.597%; 1.10: 1.633%; 1.11: 1.865%; 1.12: 1.15%; 1.13: 2.308%
| PDM3 |  | 1.01: -1.513%; 1.05: 3.436%; 1.06: 6.124%; 1.07: -0.5214%; 1.08: 2.361%; 1.09: -1.591%; 1.10: 1.086%; 1.11: -0.8634%; 1.12: -13.17%; 1.13: -7.718%
| Z-Score (Absolute) |  | 1.01: 0.23; 1.05: 1.15; 1.06: 1.9; 1.07: 0.05; 1.08: 0.85; 1.09: 0.25; 1.10: 0.5; 1.11: 0.05; 1.12: 3.48; 1.13: 1.96
| Analytical results for Ni in OREAS 238b (Certified Value 65 ppm) | [27 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 65.2 | 60.68 | 70 | 67 | 62.5333 | 65 | 74.6 | 63 | 68.3 | 66.7 | 68.2 | 67.8 | 63.6 | 64.3 | 67.7762 | 68 | 49.74 | 68 | 58.63 | 64.1 | 67.4 | 68.3 | 64.24 | 66 | 63.0439 | 63.6305 | 1
| 1 | 2 | 64.5 | 61.42 | 70 | 67 | 63.945 | 65 | 81.2 | 63 | 69.1 | 66.5 | 72 | 68.7 | 64.4 | 62.4 | 68.8781 | 66.9 | 50.5 | 66 | 60.45 | 63 | 61.8 | 67.9 | 63.25 | 69 | 61.1589 | 63.1148 | 34
| 1 | 3 | 66.1 | 61.3 | 70 | 68 | 62.185 | 66 | 71.4 | 65 | 65.5 | 64.3 | 71.1 | 66.3 | 64.1 | 63.7 | 67.5686 | 67.4 | 54.69 | 66 | 60.78 | 76.7 | 61.9 | 66.8 | 64.24 | 70 | 61.7587 | 64.7425 | 16
| 1 | 4 | 66.3 | 61.49 | 70 | 67 | 63.565 | 65 | 79.7 | 64 | 67.5 | 65.4 | 70.6 | 68.2 | 64.1 | 64 | 67.6629 | 66.8 | 55.27 | 66 | 60.78 | 64.1 | 58.3 | 68.7 | 63.69 | 71 | 56.1895 | 64.056 | 65.2221
| 1 | 5 | 66.5 | 61.27 | 70 | 67 | 62.62 | 66 | 72.3 | 62 | 65.6 | 65 | 68 | 66.4 | 65.5 | 63.7 | 69.5735 | 65.3 | 58.46 | 66 | 60.2 | 63 | 65.2 | 67.8 | 63.8 | 69 | 56.575 | 63.8692 | 47
| 1 | 6 | 66.5 | 60.15 | 75 | 69 | 62.0667 | 66 | 71.4 | 64 | 66.3 | 67 | 65 | 69.5 | 66.8 | 63.2 | 68.573 | 67.2 | 57.18 | 66 | 58.62 | 61.7 | 64 | 68.3 | 62.92 | 67 | 58.1601 | 63.3009
| Mean |  | 65.85 | 61.0517 | 70.8333 | 67.5 | 62.8192 | 65.5 | 75.1 | 63.5 | 67.05 | 65.8167 | 69.15 | 67.8167 | 64.75 | 63.55 | 68.3387 | 66.9333 | 54.3067 | 66.3333 | 59.91 | 65.4333 | 63.1 | 67.9667 | 63.69 | 68.6667 | 59.481 | 63.7857
| Median |  | 66.2 | 61.285 | 70 | 67 | 62.5767 | 65.5 | 73.45 | 63.5 | 66.9 | 65.95 | 69.4 | 68 | 64.25 | 63.7 | 68.1746 | 67.05 | 54.98 | 66 | 60.325 | 63.55 | 62.95 | 68.1 | 63.745 | 69 | 59.6595 | 63.7499
| Std Dev. |  | 0.81915 | 0.52693 | 2.0412 | 0.83666 | 0.76336 | 0.54772 | 4.3322 | 1.0488 | 1.4856 | 1.0759 | 2.5875 | 1.2703 | 1.1879 | 0.67157 | 0.80457 | 0.90701 | 3.5193 | 0.8165 | 1.019 | 5.5906 | 3.1585 | 0.65625 | 0.52983 | 1.8619 | 2.8886 | 0.58391
| Rel.Std.Dev. |  | 1.244% | 0.8631% | 2.882% | 1.239% | 1.215% | 0.8362% | 5.769% | 1.652% | 2.216% | 1.635% | 3.742% | 1.873% | 1.835% | 1.057% | 1.177% | 1.355% | 6.48% | 1.231% | 1.701% | 8.544% | 5.006% | 0.9655% | 0.8319% | 2.712% | 4.856% | 0.9154%
| PDM3 |  | 0.9628% | -6.394% | 8.603% | 3.493% | -3.684% | 0.4261% | 15.15% | -2.64% | 2.803% | 0.9117% | 6.022% | 3.978% | -0.7238% | -2.564% | 4.779% | 2.624% | -16.74% | 1.704% | -8.145% | 0.3239% | -3.254% | 4.208% | -2.349% | 5.281% | -8.802% | -2.202%
| Z-Score (Absolute) |  | 0.06 | 1.44 | 1.62 | 0.58 | 0.89 | 0.05 | 2.95 | 0.67 | 0.43 | 0.05 | 1.09 | 0.67 | 0.28 | 0.66 | 0.84 | 0.4 | 3.55 | 0.21 | 1.8 | 0.07 | 0.8 | 0.72 | 0.61 | 0.94 | 1.93 | 0.59
| Analytical results for P in OREAS 238b (Certified Value 0.065 wt.%) | [27 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 |  |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.0646 | 0.062779 | 0.06 | 0.062 | 0.062563 | 0.068 | 0.0635 | 0.064 | 0.066 | 0.066 | 0.066 | 0.067 | 0.065 | 0.0682 | 0.06517 | 0.0694 | 0.07 | 0.07 | 0.0641 | 0.0665 | 0.064 | 0.063 | 0.0665 | 0.0729 |  |  | 1
| 1 | 2 | 0.0636 | 0.062856 | 0.055 | 0.063 | 0.062254 | 0.068 | 0.0639 | 0.065 | 0.066 | 0.067 | 0.068 | 0.066 | 0.066 | 0.0656 | 0.066046 | 0.0703 | 0.07 | 0.07 | 0.0637 | 0.067 | 0.064 | 0.0629 | 0.0664 | 0.0728 |  |  | 0
| 1 | 3 | 0.0632 | 0.062712 | 0.065 | 0.063 | 0.062499 | 0.068 | 0.0646 | 0.065 | 0.061 | 0.067 | 0.069 | 0.067 | 0.063 | 0.0663 | 0.06476 | 0.0729 | 0.07 | 0.07 | 0.0644 | 0.0672 | 0.065 | 0.0645 | 0.0661 | 0.0739 |  |  | 16
| 1 | 4 | 0.0653 | 0.062939 | 0.06 | 0.062 | 0.063137 | 0.069 | 0.0655 | 0.065 | 0.061 | 0.069 | 0.066 | 0.068 | 0.067 | 0.0671 | 0.064833 | 0.0689 | 0.06 | 0.07 | 0.0639 | 0.0672 | 0.065 | 0.0645 | 0.0688 | 0.0703 |  |  | 0.065288
| 1 | 5 | 0.0636 | 0.062999 | 0.06 | 0.063 | 0.063341 | 0.069 | 0.0648 | 0.064 | 0.061 | 0.067 | 0.066 | 0.066 | 0.069 | 0.066 | 0.064848 | 0.0663 | 0.06 | 0.07 | 0.0623 | 0.067 | 0.065 | 0.0649 | 0.0656 | 0.0712 |  |  | 48
| 1 | 6 | 0.0647 | 0.06215 | 0.06 | 0.063 | 0.063191 | 0.068 | 0.0626 | 0.064 | 0.062 | 0.069 | 0.066 | 0.067 | 0.067 | 0.065 | 0.063863 | 0.0689 | 0.06 | 0.07 | 0.0643 | 0.0678 | 0.066 | 0.0627 | 0.0643 | 0.071
| Mean |  | 0.064167 | 0.062739 | 0.06 | 0.062667 | 0.062831 | 0.068333 | 0.06415 | 0.0645 | 0.062833 | 0.0675 | 0.066833 | 0.066833 | 0.066167 | 0.066367 | 0.06492 | 0.06945 | 0.065 | 0.07 | 0.063783 | 0.067117 | 0.064833 | 0.06375 | 0.066283 | 0.072017
| Median |  | 0.0641 | 0.062818 | 0.06 | 0.063 | 0.06285 | 0.068 | 0.06425 | 0.0645 | 0.0615 | 0.067 | 0.066 | 0.067 | 0.0665 | 0.06615 | 0.064841 | 0.06915 | 0.065 | 0.07 | 0.064 | 0.0671 | 0.065 | 0.06375 | 0.06625 | 0.072
| Std Dev. |  | 0.0008165 | 0.00030661 | 0.0031623 | 0.0005164 | 0.0004468 | 0.0005164 | 0.001033 | 0.00054772 | 0.0024833 | 0.0012247 | 0.0013292 | 0.00075277 | 0.0020412 | 0.0011396 | 0.00070529 | 0.002152 | 0.0054772 | 0 | 0.0007705 | 0.00042151 | 0.00075277 | 0.00098336 | 0.0014716 | 0.0013848
| Rel.Std.Dev. |  | 1.272% | 0.4887% | 5.27% | 0.824% | 0.7111% | 0.7557% | 1.61% | 0.8492% | 3.952% | 1.814% | 1.989% | 1.126% | 3.085% | 1.717% | 1.086% | 3.099% | 8.427% | 0% | 1.208% | 0.628% | 1.161% | 1.543% | 2.22% | 1.923%
| PDM3 |  | -1.718% | -3.904% | -8.1% | -4.015% | -3.764% | 4.664% | -1.743% | -1.207% | -3.76% | 3.388% | 2.367% | 2.367% | 1.345% | 1.652% | -0.564% | 6.374% | -0.4415% | 7.217% | -2.305% | 2.801% | -0.6968% | -2.356% | 1.524% | 10.31%
| Z-Score (Absolute) |  | 0.29 | 0.8 | 1.79 | 0.83 | 0.77 | 1.21 | 0.29 | 0.17 | 0.77 | 0.91 | 0.67 | 0.67 | 0.43 | 0.51 | 0.01 | 1.62 | 0.01 | 1.81 | 0.42 | 0.78 | 0.05 | 0.44 | 0.48 | 2.54
| Analytical results for Pb in OREAS 238b (Certified Value 23.4 ppm) | [27 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 25 | 24 | 21 | 23.708 | 23 | 21.1 | 24 | 19.3 | 24.2 | 22.1 | 24.5 | 23.6 | 22.65 | 25.5989 | 22.7 | 23.55 | 22.7 | 23.09 | 25.5 | 23.7 | 23.05 | 22 | 20 | 26.4605 | 24.19 |  | 1
| 1 | 2 | 24.3 | 23 | 23 | 23.2787 | 24 | 21.8 | 24 | 20.7 | 24.1 | 23.9 | 24.2 | 24 | 22.61 | 25.1396 | 23.8 | 23.54 | 24.7 | 22.82 | 25.9 | 21.2 | 23.67 | 19.6 | 21 | 27.2067 | 22.828 |  | 35
| 1 | 3 | 24.6 | 23 | 21 | 21.9633 | 24 | 20.4 | 27 | 18 | 23.9 | 24.4 | 23.8 | 23.1 | 23.43 | 24.5114 | 24.8 | 23.31 | 23.4 | 23.36 | 26.7 | 21.5 | 23.14 | 21.2 | 20 | 26.2118 | 23.6898 |  | 16
| 1 | 4 | 24.8 | 26 | 23 | 21.5 | 24 | 20.2 | 24 | 19.9 | 24.7 | 23 | 23.6 | 23.7 | 22.81 | 24.4991 | 24 | 22.44 | 23.1 | 22.88 | 25.4 | 19.7 | 24.34 | 21 | 22 | 23.1774 | 23.4536 |  | 23.4425
| 1 | 5 | 24.6 | 23 | 19 | 21.3807 | 24 | 22.6 | 24 | 19.7 | 24.4 | 23.2 | 22.9 | 25.3 | 23.58 | 24.7212 | 24 | 24.3 | 23.4 | 21.97 | 24.3 | 22.4 | 22.99 | 22.6 | 20 | 23.7246 | 24.1554 |  | 49
| 1 | 6 | 24.3 | 24 | 22 | 23.048 | 23 | 21.9 | 24 | 19.3 | 24.7 | 24.3 | 24.6 | 24.1 | 24.21 | 24.8811 | 23.2 | 23.98 | 22.8 | 22.5 | 24.8 | 21.9 | 22.85 | 20.8 | 19 | 23.0779 | 23.0925
| Mean |  | 24.6 | 23.8333 | 21.5 | 22.4798 | 23.6667 | 21.3333 | 24.5 | 19.4833 | 24.3333 | 23.4833 | 23.9333 | 23.9667 | 23.215 | 24.8919 | 23.75 | 23.52 | 23.35 | 22.77 | 25.4333 | 21.7333 | 23.34 | 21.2 | 20.3333 | 24.9765 | 23.5682
| Median |  | 24.6 | 23.5 | 21.5 | 22.5057 | 24 | 21.45 | 24 | 19.5 | 24.3 | 23.55 | 24 | 23.85 | 23.12 | 24.8011 | 23.9 | 23.545 | 23.25 | 22.85 | 25.45 | 21.7 | 23.095 | 21.1 | 20 | 24.9682 | 23.5717
| Std Dev. |  | 0.27568 | 1.169 | 1.5166 | 0.99039 | 0.5164 | 0.9331 | 1.2247 | 0.89088 | 0.3266 | 0.88412 | 0.6377 | 0.74207 | 0.63541 | 0.42181 | 0.72595 | 0.63746 | 0.72319 | 0.48539 | 0.83825 | 1.3277 | 0.56491 | 1.0354 | 1.0328 | 1.8499 | 0.55382
| Rel.Std.Dev. |  | 1.121% | 4.905% | 7.054% | 4.406% | 2.182% | 4.374% | 4.999% | 4.573% | 1.342% | 3.765% | 2.665% | 3.096% | 2.737% | 1.695% | 3.057% | 2.71% | 3.097% | 2.132% | 3.296% | 6.109% | 2.42% | 4.884% | 5.079% | 7.407% | 2.35%
| PDM3 |  | 4.938% | 1.667% | -8.286% | -4.107% | 0.9561% | -8.997% | 4.511% | -16.89% | 3.8% | 0.1741% | 2.094% | 2.236% | -0.9706% | 6.183% | 1.312% | 0.3305% | -0.3947% | -2.869% | 8.492% | -7.291% | -0.4373% | -9.566% | -13.26% | 6.543% | 0.5362%
| Z-Score (Absolute) |  | 0.9 | 0.26 | 1.67 | 0.86 | 0.12 | 1.81 | 0.81 | 3.35 | 0.67 | 0.03 | 0.34 | 0.37 | 0.25 | 1.14 | 0.19 | 0 | 0.14 | 0.62 | 1.59 | 1.48 | 0.15 | 1.92 | 2.64 | 1.21 | 0.04
| Analytical results for Pd in OREAS 238b (Indicative Value 5.01 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.30 |  | ppb
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: < 7.5 |  | 1
| 1 | 2 | 1.01: 9.7 |  | 1
| 1 | 3 | 1.01: < 5.3 |  | 16
| 1 | 4 | 1.01: < 3.2 |  | 5.0083
| 1 | 5 | 1.01: < 11.7 |  | 10
| 1 | 6 | 1.01: 6.5
| Mean |  | 1.01: 8.1
| Median |  | 1.01: 8.1
| Std Dev. |  | 1.01: 2.2627
| Rel.Std.Dev. |  | 1.01: 27.94%
| PDM3 |  | 1.01: 61.73%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Pr in OREAS 238b (Certified Value 9.31 ppm) | [27 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.03; 1.05: 1.05; 1.06: 1.06; 1.07: 1.16; 1.08: 1.19; 1.09: 1.20; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*MS; 1.13: 4A*MS |  | 2
3
| 1 | 1 | 1.01: 8.8826; 1.05: 9.45; 1.06: 9.5; 1.07: 9.6; 1.08: 10.3; 1.09: 8.946; 1.10: 9.48; 1.11: 9.1; 1.12: 7.5; 1.13: 8.7309 |  | 1
| 1 | 2 | 1.01: 8.8187; 1.05: 9.4; 1.06: 9.8; 1.07: 8.2; 1.08: 9.8; 1.09: 8.886; 1.10: 9.62; 1.11: 9.6; 1.12: 7.18; 1.13: 8.2913 |  | 11
| 1 | 3 | 1.01: 8.8145; 1.05: 9.55; 1.06: 9.7; 1.07: 11.3; 1.08: 10.1; 1.09: 9.378; 1.10: 9.72; 1.11: 9.5; 1.12: 7.57; 1.13: 8.3963 |  | 16
| 1 | 4 | 1.01: 8.9045; 1.05: 9.9; 1.06: 9.7; 1.07: 9.3; 1.08: 9.7; 1.09: 8.947; 1.10: 9.66; 1.11: 9.6; 1.12: 7.41; 1.13: 8.5503 |  | 9.3058
| 1 | 5 | 1.01: 8.7119; 1.05: 9.9; 1.06: 9.7; 1.07: 9.8; 1.08: 9.6; 1.09: 9.208; 1.10: 9.47; 1.11: 9.1; 1.12: 7.38; 1.13: 8.3011 |  | 50
| 1 | 6 | 1.01: 8.9223; 1.05: 9.9; 1.06: 9.6; 1.07: 9.4; 1.08: 9.8; 1.09: 9.108; 1.10: 9.43; 1.11: 9.1; 1.12: 7.46; 1.13: 8.3754
| Mean |  | 1.01: 8.8424; 1.05: 9.6833; 1.06: 9.6667; 1.07: 9.6; 1.08: 9.8833; 1.09: 9.0788; 1.10: 9.5633; 1.11: 9.3333; 1.12: 7.4167; 1.13: 8.4409
| Median |  | 1.01: 8.8507; 1.05: 9.725; 1.06: 9.7; 1.07: 9.5; 1.08: 9.8; 1.09: 9.0275; 1.10: 9.55; 1.11: 9.3; 1.12: 7.435; 1.13: 8.3859
| Std Dev. |  | 1.01: 0.07777; 1.05: 0.24221; 1.06: 0.10328; 1.07: 1.002; 1.08: 0.26394; 1.09: 0.18934; 1.10: 0.11877; 1.11: 0.2582; 1.12: 0.13397; 1.13: 0.16987
| Rel.Std.Dev. |  | 1.01: 0.8795%; 1.05: 2.501%; 1.06: 1.068%; 1.07: 10.44%; 1.08: 2.671%; 1.09: 2.085%; 1.10: 1.242%; 1.11: 2.766%; 1.12: 1.806%; 1.13: 2.012%
| PDM3 |  | 1.01: -4.979%; 1.05: 4.057%; 1.06: 3.878%; 1.07: 3.162%; 1.08: 6.206%; 1.09: -2.439%; 1.10: 2.768%; 1.11: 0.2959%; 1.12: -20.3%; 1.13: -9.294%
| Z-Score (Absolute) |  | 1.01: 1.35; 1.05: 0.52; 1.06: 0.49; 1.07: 0.34; 1.08: 0.97; 1.09: 0.82; 1.10: 0.26; 1.11: 0.26; 1.12: 4.53; 1.13: 2.25
| Analytical results for Pt in OREAS 238b (Indicative Value 1.09 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.30 |  | ppb
|  |  | 1.01: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.2 |  | 1
| 1 | 2 | 1.01: < 3.8 |  | 1
| 1 | 3 | 1.01: 2.3 |  | 16
| 1 | 4 | 1.01: < 2.8 |  | 1.0917
| 1 | 5 | 1.01: 0.7 |  | 11
| 1 | 6 | 1.01: < 0.1
| Mean |  | 1.01: 1.0667
| Median |  | 1.01: 0.7
| Std Dev. |  | 1.01: 1.097
| Rel.Std.Dev. |  | 1.01: 102.8%
| PDM3 |  | 1.01: -2.29%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Rb in OREAS 238b (Certified Value 159 ppm) | [27 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 162.42 | 143.3728 | 172 | 171 | 166.4 | 168.6 | 153 | 159.5 | 160.5 | 160.5 | 162 | 163.5 | 235.3 | 152.1169 | 162 | 145.6 | 174 | 157.69 | 154 | 155.4 | 159.861 | 131.9752 |  |  |  |  | 1
| 1 | 2 | 161.78 | 143.7192 | 166 | 176 | 163 | 170.2 | 152 | 160.5 | 162.5 | 159.5 | 153 | 163.5 | 166 | 150.026 | 153 | 143.6 | 166 | 139.96 | 140.4 | 156.5 | 163.5221 | 135.9644 |  |  |  |  | 37
| 1 | 3 | 165.81 | 143.8721 | 165 | 181 | 165.9 | 169.6 | 150 | 158.8 | 154.5 | 160 | 156.5 | 164.5 | 247.7 | 147.2059 | 158 | 148.3 | 162 | 153.03 | 140.5 | 155 | 162.4864 | 125.6283 |  |  |  |  | 16
| 1 | 4 | 168.79 | 141.5726 | 173 | 184 | 164.5 | 174.9 | 158 | 158.4 | 164.5 | 159 | 161.5 | 165.5 | 202.2 | 150.2705 | 157 | 146.2 | 169 | 149.82 | 132.1 | 156.8 | 161.596 | 122.7831 |  |  |  |  | 158.6634
| 1 | 5 | 166.7 | 141.2279 | 168 | 176 | 163.9 | 171.6 | 119 | 158.9 | 156 | 160 | 151 | 162.5 | 211.8 | 144.0567 | 153 | 152.7 | 163 | 160.66 | 147.6 | 156 | 161.0135 | 114.4306 |  |  |  |  | 51
| 1 | 6 | 168.29 | 141.1467 | 170 | 182 | 162.7 | 172.2 | 152 | 158 | 163 | 165.5 | 158.5 | 164.5 | 170.6 | 145.6879 | 152 | 144.8 | 159 | 159.39 | 146.6 | 156 | 157.5167 | 134.1775
| Mean |  | 165.6317 | 142.4852 | 169 | 178.3333 | 164.4 | 171.1833 | 147.3333 | 159.0167 | 160.1667 | 160.75 | 157.0833 | 164 | 205.6 | 148.2273 | 155.8333 | 146.8667 | 165.5 | 153.425 | 143.5333 | 155.95 | 160.9993 | 127.4932
| Median |  | 166.255 | 142.4727 | 169 | 178.5 | 164.2 | 170.9 | 152 | 158.85 | 161.5 | 160 | 157.5 | 164 | 207 | 148.6159 | 155 | 145.9 | 164.5 | 155.36 | 143.55 | 156 | 161.3048 | 128.8017
| Std Dev. |  | 2.945 | 1.2992 | 3.2249 | 4.8442 | 1.5073 | 2.2436 | 14.1374 | 0.88412 | 4.0456 | 2.3822 | 4.4656 | 1.0488 | 33.1592 | 3.0798 | 3.8687 | 3.258 | 5.3944 | 7.7534 | 7.5529 | 0.66858 | 2.1145 | 8.1587
| Rel.Std.Dev. |  | 1.778% | 0.9118% | 1.908% | 2.716% | 0.9169% | 1.311% | 9.596% | 0.556% | 2.526% | 1.482% | 2.843% | 0.6395% | 16.13% | 2.078% | 2.483% | 2.218% | 3.259% | 5.054% | 5.262% | 0.4287% | 1.313% | 6.399%
| PDM3 |  | 4.392% | -10.2% | 6.515% | 12.4% | 3.616% | 7.891% | -7.141% | 0.2226% | 0.9474% | 1.315% | -0.9959% | 3.363% | 29.58% | -6.578% | -1.784% | -7.435% | 4.309% | -3.302% | -9.536% | -1.71% | 1.472% | -19.65%
| Z-Score (Absolute) |  | 0.67 | 1.89 | 1.04 | 2.07 | 0.53 | 1.28 | 1.35 | 0.06 | 0.06 | 0.13 | 0.28 | 0.49 | 5.08 | 1.26 | 0.42 | 1.41 | 0.65 | 0.68 | 1.77 | 0.4 | 0.16 | 3.55
| Analytical results for Re in OREAS 238b (Certified Value < 0.002 ppm) | [27 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.26 | 1.27 | 1.28 | 1.30 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.002 | 0.0003 | [10 empty] | 1
| 1 | 2 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.006 | 0.001 | < 0.002 | < 0.002 | 0.0012 | [10 empty] | 38
| 1 | 3 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.005 | < 0.001 | < 0.002 | < 0.002 | 0.0009 | [10 empty] | 16
| 1 | 4 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.003 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.006 | 0.001 | < 0.002 | < 0.002 | 0.0006 | [10 empty] | < 0.002
| 1 | 5 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.007 | < 0.001 | < 0.002 | < 0.002 | 0.0015 | [10 empty] | 52
| 1 | 6 | < 0.002 | < 0.1 | < 0.001 | < 0.05 | 0.004 | < 0.05 | < 0.005 | < 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.005 | 0.001 | < 0.002 | < 0.002 | 0.0008
| Mean |  | 1.09: 0.0036667; 1.14: 0.002; 1.16: 0.0058; 1.26: 0.0015; 1.30: 0.00088333
| Median |  | 1.09: 0.004; 1.14: 0.002; 1.16: 0.006; 1.26: 0.001; 1.30: 0.00085
| Std Dev. |  | 1.09: 0.0005164; 1.14: 0; 1.16: 0.00083666; 1.26: 0.001; 1.30: 0.00042622
| Rel.Std.Dev. |  | 1.09: 14.08%; 1.14: 0%; 1.16: 14.43%; 1.26: 66.67%; 1.30: 48.25%
PDM3
| Z-Score (Absolute) |  | 0.14 | 84.22 | 1 | 41.18 | 4.45 | 41.18 | 2.44 | 0.14 | 0.14 | 1 | 0.14 | 6.74 | 0.14 | 0.14 | 0.14 | 0.34
| Analytical results for Rh in OREAS 238b (Indicative Value 0.001 ppm) | [27 empty] | Indicative
Rh | Round | Replicate | 1.01: Lab | [10 empty] | 1
| No | No | 1.01: 1.30 | [10 empty] | ppm
|  |  | 1.01: 4A*MS | [10 empty] | 3
3
| 1 | 1 | 1.01: 0.0005 | [10 empty] | 1
| 1 | 2 | 1.01: 0.0007 | [10 empty] | 6
| 1 | 3 | 1.01: 0.0007 | [10 empty] | 16
| 1 | 4 | 1.01: 0.0005 | [10 empty] | 0.00081667
| 1 | 5 | 1.01: 0.0014 | [10 empty] | 12
| 1 | 6 | 1.01: 0.0011
| Mean |  | 1.01: 0.00081667
| Median |  | 1.01: 0.0007
| Std Dev. |  | 1.01: 0.00036009
| Rel.Std.Dev. |  | 1.01: 44.09%
| PDM3 |  | 1.01: -5.551e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ru in OREAS 238b (Indicative Value 0.003 ppm) | [27 empty] | Indicative
Ru | Round | Replicate | 1.01: Lab | [10 empty] | 1
| No | No | 1.01: 1.30 | [10 empty] | ppm
|  |  | 1.01: 4A*MS | [10 empty] | 3
3
| 1 | 1 | 1.01: 0.0037 | [10 empty] | 1
| 1 | 2 | 1.01: < 0.0012 | [10 empty] | 14
| 1 | 3 | 1.01: 0.0045 | [10 empty] | 16
| 1 | 4 | 1.01: 0.0029 | [10 empty] | 0.0033167
| 1 | 5 | 1.01: 0.0061 | [10 empty] | 9
| 1 | 6 | 1.01: 0.0021
| Mean |  | 1.01: 0.00386
| Median |  | 1.01: 0.0037
| Std Dev. |  | 1.01: 0.0015388
| Rel.Std.Dev. |  | 1.01: 39.87%
| PDM3 |  | 1.01: 16.38%
Z-Score (Absolute) | bar: 1.01
| Analytical results for S in OREAS 238b (Certified Value 0.175 wt.%) | [27 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [5 empty] | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.19 | 0.18 | 0.17 | 0.18 | 0.179 | 0.13 | 0.177 | 0.17 | 0.17 | 0.18 | 0.17 | 0.18 | 0.1906 | 0.177 | 0.18 | 0.17 | 0.17 | 0.18 | 0.17 | 0.16169 | 0.1806 | [5 empty] | 1
| 1 | 2 | 0.18 | 0.175 | 0.17 | 0.165 | 0.179 | 0.15 | 0.181 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.1782 | 0.18 | 0.18 | 0.17 | 0.17 | 0.18 | 0.17 | 0.16254 | 0.1764 | [5 empty] | 23
| 1 | 3 | 0.18 | 0.17 | 0.17 | 0.165 | 0.183 | 0.15 | 0.179 | 0.17 | 0.17 | 0.19 | 0.17 | 0.17 | 0.1947 | 0.18 | 0.17 | 0.17 | 0.17 | 0.18 | 0.17 | 0.16559 | 0.1877 | [5 empty] | 16
| 1 | 4 | 0.17 | 0.18 | 0.17 | 0.17 | 0.183 | 0.16 | 0.178 | 0.18 | 0.18 | 0.18 | 0.17 | 0.17 | 0.1936 | 0.175 | 0.18 | 0.16 | 0.17 | 0.18 | 0.17 | 0.16426 | 0.1674 | [5 empty] | 0.17477
| 1 | 5 | 0.16 | 0.18 | 0.17 | 0.17 | 0.18 | 0.15 | 0.172 | 0.17 | 0.18 | 0.18 | 0.17 | 0.18 | 0.1895 | 0.178 | 0.18 | 0.17 | 0.17 | 0.18 | 0.17 | 0.16587 | 0.1622 | [5 empty] | 53
| 1 | 6 | 0.2 | 0.175 | 0.17 | 0.18 | 0.18 | 0.15 | 0.177 | 0.17 | 0.18 | 0.18 | 0.17 | 0.18 | 0.1741 | 0.177 | 0.18 | 0.16 | 0.18 | 0.18 | 0.17 | 0.16302 | 0.1637
| Mean |  | 0.18 | 0.17667 | 0.17 | 0.17167 | 0.18067 | 0.14833 | 0.17733 | 0.17 | 0.175 | 0.18333 | 0.17 | 0.17667 | 0.18678 | 0.17783 | 0.17833 | 0.16667 | 0.17167 | 0.18 | 0.17 | 0.16383 | 0.173
| Median |  | 0.18 | 0.1775 | 0.17 | 0.17 | 0.18 | 0.15 | 0.1775 | 0.17 | 0.175 | 0.18 | 0.17 | 0.18 | 0.19005 | 0.1775 | 0.18 | 0.17 | 0.17 | 0.18 | 0.17 | 0.16364 | 0.1719
| Std Dev. |  | 0.014142 | 0.0040825 | 0 | 0.0068313 | 0.0018619 | 0.0098319 | 0.0030111 | 0.0063246 | 0.0054772 | 0.005164 | 0 | 0.005164 | 0.0085514 | 0.0019408 | 0.0040825 | 0.005164 | 0.0040825 | 3.0405e-17 | 0 | 0.0016917 | 0.010197
| Rel.Std.Dev. |  | 7.857% | 2.311% | 0% | 3.979% | 1.031% | 6.628% | 1.698% | 3.72% | 3.13% | 2.817% | 0% | 2.923% | 4.578% | 1.091% | 2.289% | 3.098% | 2.378% | 1.689e-14% | 0% | 1.033% | 5.894%
| PDM3 |  | 2.991% | 1.084% | -2.731% | -1.777% | 3.373% | -15.13% | 1.465% | -2.731% | 0.1303% | 4.898% | -2.731% | 1.084% | 6.872% | 1.751% | 2.038% | -4.638% | -1.777% | 2.991% | -2.731% | -6.262% | -1.014%
| Z-Score (Absolute) |  | 0.67 | 0.22 | 0.67 | 0.45 | 0.76 | 3.6 | 0.31 | 0.67 | 0 | 1.12 | 0.67 | 0.22 | 1.59 | 0.38 | 0.45 | 1.12 | 0.45 | 0.67 | 0.67 | 1.51 | 0.27
| Analytical results for Sb in OREAS 238b (Certified Value 620 ppm) | [27 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  |  |  | 0
0
| 1 | 1 | 670.22 | 15.5 | 609.9556 | 651.4 | 628.38 | 568 | 583.5 | 643 | 707 | 637 | 637 | 633.18 | 641.9401 | 512.7 | 600 | 583.3 | 616.69 | 596.2 | 605.68 | 630.979 | 696 | 315.5259 | 642.119 |  |  |  | 1
| 1 | 2 | 668.26 | 21.3 | 611.4381 | 657.5 | 611.06 | 583 | 598.5 | 646 | 716 | 619 | 620 | 620.55 | 626.7426 | 506.3 | 606 | 587.3 | 602.46 | 603.1 | 629.21 | 610.863 | 699 | 324.4039 | 654.424 |  |  |  | 39
| 1 | 3 | 681.44 | 17.7 | 616.92 | 661.6 | 603.98 | 651 | 582.1 | 639 | 723 | 627 | 617 | 632.89 | 638.7533 | 524.4 | 587 | 587.7 | 617.21 | 612.3 | 616.23 | 632.584 | 693 | 328.7533 | 632.3607 |  |  |  | 16
| 1 | 4 | 680.68 | 16.1 | 610.3533 | 661.1 | 607.66 | 589 | 587.6 | 663 | 697 | 621 | 625 | 637.26 | 622.8416 | 514.2 | 595 | 583.5 | 617.14 | 605.4 | 633.53 | 611.826 | 728 | 311.4835 | 641.3478 |  |  |  | 620.187
| 1 | 5 | 681.13 | 14.3 | 618.18 | 655.2 | 614.67 | 561 | 585 | 649 | 692 | 632 | 643 | 632.21 | 634.539 | 543.6 | 597 | 592.7 | 604.87 | 598 | 609.02 | 628.625 | 684 | 315.0142 | 614.2106 |  |  |  | 54
| 1 | 6 | 681.88 | 19.4 | 612.6978 | 653.7 | 633.06 | 558 | 586 | 661 | 604 | 638 | 643 | 626.15 | 631.8134 | 511.1 | 592 | 585.1 | 581.89 | 602.2 | 604.24 | 622.419 | 672 | 315.1421 | 631.6317
| Mean |  | 677.2683 | 17.3833 | 613.2575 | 656.75 | 616.4683 | 585 | 587.1167 | 650.1667 | 689.8333 | 629 | 630.8333 | 630.3733 | 632.7717 | 518.7167 | 596.1667 | 586.6 | 606.71 | 602.8667 | 616.3183 | 622.8827 | 695.3333 | 318.3872 | 636.0156
| Median |  | 680.905 | 16.9 | 612.068 | 656.35 | 612.865 | 575.5 | 585.5 | 647.5 | 702 | 629.5 | 631 | 632.55 | 633.1762 | 513.45 | 596 | 586.2 | 610.78 | 602.65 | 612.625 | 625.522 | 694.5 | 315.334 | 636.8542
| Std Dev. |  | 6.2618 | 2.6157 | 3.4815 | 4.0825 | 11.6893 | 34.5543 | 5.8956 | 9.7656 | 43.5955 | 8.025 | 11.6347 | 5.9891 | 7.195 | 13.5666 | 6.5549 | 3.5117 | 13.8283 | 5.7277 | 12.4481 | 9.5873 | 18.7581 | 6.6545 | 13.5099
| Rel.Std.Dev. |  | 0.9246% | 15.05% | 0.5677% | 0.6216% | 1.896% | 5.907% | 1.004% | 1.502% | 6.32% | 1.276% | 1.844% | 0.9501% | 1.137% | 2.615% | 1.1% | 0.5987% | 2.279% | 0.9501% | 2.02% | 1.539% | 2.698% | 2.09% | 2.124%
| PDM3 |  | 9.204% | -97.2% | -1.117% | 5.895% | -0.5996% | -5.674% | -5.332% | 4.834% | 11.23% | 1.421% | 1.717% | 1.642% | 2.029% | -16.36% | -3.873% | -5.416% | -2.173% | -2.793% | -0.6238% | 0.4347% | 12.12% | -48.66% | 2.552%
| Z-Score (Absolute) |  | 2.02 | 19.9 | 0.11 | 1.34 | 0 | 1.05 | 0.97 | 1.12 | 2.44 | 0.42 | 0.48 | 0.46 | 0.54 | 3.25 | 0.67 | 0.99 | 0.32 | 0.45 | 0 | 0.21 | 2.62 | 9.9 | 0.65
| Analytical results for Sc in OREAS 238b (Certified Value 14.3 ppm) | [27 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.6 | 13.7685 | 16 | 14 | 10.8587 | 14 | 13.5 | 14 | 14.6 | 15.4 | 15.8 | 15.4 | 15 | 13.9 | 14.2355 | 14.6 | 12.53 | 14.6 | 14.2 | 14.3 | 14.9 | 14.4375 | 15.2273 | 10.3198 |  |  | 1
| 1 | 2 | 14.5 | 13.5088 | 16 | 14 | 11.1807 | 14 | 12 | 14 | 14.5 | 15.7 | 16 | 15.1 | 15.1 | 12.8 | 14.3704 | 13 | 12.41 | 14.6 | 14.1 | 12.4 | 14.5 | 14.238 | 15.2844 | 10.1249 |  |  | 40
| 1 | 3 | 14.4 | 13.7429 | 16 | 14 | 10.251 | 15 | 13 | 14 | 13.4 | 14.7 | 16.2 | 14.8 | 15.1 | 12.7 | 14.4406 | 13.7 | 12.47 | 13.4 | 14.7 | 12.2 | 14.7 | 14.6055 | 15.5701 | 10.584 |  |  | 16
| 1 | 4 | 14.7 | 13.9419 | 17 | 15 | 10.3073 | 15 | 14.4 | 14 | 13.6 | 15.2 | 15.4 | 15.5 | 15.1 | 13.1 | 14.2993 | 13.7 | 12.89 | 13.7 | 14.5 | 11.4 | 14.9 | 14.4585 | 14.7702 | 9.93 |  |  | 14.2704
| 1 | 5 | 15 | 13.4984 | 16 | 14 | 10.746 | 14 | 13.8 | 14 | 13.2 | 14.8 | 15.5 | 14.6 | 15.3 | 12.7 | 14.1979 | 13.3 | 12.25 | 14.3 | 13.8 | 12.9 | 14.6 | 14.5005 | 14.9987 | 9.3628 |  |  | 55
| 1 | 6 | 14.9 | 13.5576 | 17 | 14 | 12.1493 | 15 | 13.4 | 14 | 13.4 | 15.5 | 15.2 | 15.6 | 15.4 | 12.5 | 14.4051 | 13.7 | 12.4 | 14 | 13.4 | 12.6 | 14.5 | 14.217 | 14.8845 | 10.534
| Mean |  | 14.6833 | 13.6697 | 16.3333 | 14.1667 | 10.9155 | 14.5 | 13.35 | 14 | 13.7833 | 15.2167 | 15.6833 | 15.1667 | 15.1667 | 12.95 | 14.3248 | 13.6667 | 12.4917 | 14.1 | 14.1167 | 12.6333 | 14.6833 | 14.4095 | 15.1225 | 10.1426
| Median |  | 14.65 | 13.6502 | 16 | 14 | 10.8023 | 14.5 | 13.45 | 14 | 13.5 | 15.3 | 15.65 | 15.25 | 15.1 | 12.75 | 14.3349 | 13.7 | 12.44 | 14.15 | 14.15 | 12.5 | 14.65 | 14.448 | 15.113 | 10.2224
| Std Dev. |  | 0.23166 | 0.1772 | 0.5164 | 0.40825 | 0.698 | 0.54772 | 0.80932 | 0 | 0.608 | 0.39707 | 0.38166 | 0.40332 | 0.15055 | 0.50498 | 0.096606 | 0.53914 | 0.21637 | 0.4899 | 0.47081 | 0.96056 | 0.18348 | 0.15255 | 0.29449 | 0.45451
| Rel.Std.Dev. |  | 1.578% | 1.296% | 3.162% | 2.882% | 6.395% | 3.777% | 6.062% | 0% | 4.411% | 2.609% | 2.434% | 2.659% | 0.9927% | 3.899% | 0.6744% | 3.945% | 1.732% | 3.474% | 3.335% | 7.603% | 1.25% | 1.059% | 1.947% | 4.481%
| PDM3 |  | 2.894% | -4.21% | 14.46% | -0.7268% | -23.51% | 1.609% | -6.45% | -1.895% | -3.413% | 6.631% | 9.901% | 6.281% | 6.281% | -9.253% | 0.3815% | -4.231% | -12.46% | -1.194% | -1.077% | -11.47% | 2.894% | 0.9748% | 5.971% | -28.93%
| Z-Score (Absolute) |  | 0.55 | 0.48 | 2.22 | 0.03 | 3.26 | 0.36 | 0.8 | 0.14 | 0.36 | 1.09 | 1.56 | 1.04 | 1.04 | 1.21 | 0.19 | 0.48 | 1.67 | 0.04 | 0.03 | 1.53 | 0.55 | 0.27 | 0.99 | 4.04
| Analytical results for Se in OREAS 238b (Indicative Value 0.8 ppm) | [27 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | < 0.5 | < 5 | 0.8 | < 5 | < 2 | < 1 | 1 | < 1 | 1 | < 1 | < 1 | 1.1 | < 0.1 | < 2 | < 0.5 | < 1 | 0.4 | 0.2961 | 3.5119 | [7 empty] | 1
| 1 | 2 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 0.7 | < 0.1 | < 2 | 0.7 | < 1 | 0.3 | 0.2283 | 3.2826 | [7 empty] | 41
| 1 | 3 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 1 | < 0.1 | < 2 | < 0.5 | < 1 | 0.5 | 0.2152 | 3.3068 | [7 empty] | 16
| 1 | 4 | < 0.5 | < 5 | 1.1 | < 5 | < 2 | < 1 | 2 | < 1 | 1 | < 1 | < 1 | 1.2 | < 0.1 | < 2 | 1.7 | < 1 | 0.6 | 0.2128 | 3.2223 | [7 empty] | 0.8
| 1 | 5 | < 0.5 | < 5 | 1 | < 5 | < 2 | < 1 | < 1 | < 1 | 1 | < 1 | < 1 | 1.3 | < 0.1 | < 2 | 1.3 | < 1 | 0.5 | 0.1938 | 3.4154 | [7 empty] | 10
| 1 | 6 | < 0.5 | < 5 | 1 | < 5 | < 2 | < 1 | 1 | < 1 | < 1 | < 1 | 1 | 1.2 | < 0.1 | < 2 | 0.8 | < 1 | 0.5 | 0.2295 | 3.2464
| Mean |  | 1.07: 1.0167; 1.11: 1.3333; 1.13: 1; 1.15: 1; 1.16: 1.0833; 1.21: 1.125; 1.26: 0.46667; 1.27: 0.22928; 1.28: 3.3309
| Median |  | 1.07: 1.05; 1.11: 1; 1.13: 1; 1.15: 1; 1.16: 1.15; 1.21: 1.05; 1.26: 0.5; 1.27: 0.22175; 1.28: 3.2947
| Std Dev. |  | 1.07: 0.1169; 1.11: 0.57735; 1.13: 0; 1.16: 0.2137; 1.21: 0.46458; 1.26: 0.10328; 1.27: 0.035198; 1.28: 0.11113
| Rel.Std.Dev. |  | 1.07: 11.5%; 1.11: 43.3%; 1.13: 0%; 1.16: 19.73%; 1.21: 41.3%; 1.26: 22.13%; 1.27: 15.35%; 1.28: 3.336%
| PDM3 |  | 1.07: 27.08%; 1.11: 66.67%; 1.13: 25%; 1.15: 25%; 1.16: 35.42%; 1.21: 40.62%; 1.26: -41.67%; 1.27: -71.34%; 1.28: 316.4%
| Z-Score (Absolute) |  | 1.18 | 3.37 | 0.37 | 3.37 | 0.34 | 0.67 | 0.17 | 0.67 | 0.17 | 0.67 | 0.51 | 0.51 | 1.58 | 0.34 | 0 | 0.67 | 0.74 | 1.22 | 5.05
| Analytical results for Sm in OREAS 238b (Certified Value 6.54 ppm) | [27 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.05; 1.07: 1.06; 1.08: 1.16; 1.09: 1.19; 1.10: 1.20; 1.11: 1.26; 1.12: 1.28; 1.13: 1.30; 1.14: 1.34 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS; 1.12: 4A*OES/MS; 1.13: 4A*MS; 1.14: 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.01: 6.2977; 1.06: 6.8; 1.07: 7.1; 1.08: 6.2; 1.09: 7.2; 1.10: 6.163; 1.11: 6.75; 1.12: 6.2; 1.13: 7.4918; 1.14: 6.2299 |  |  |  | 1
| 1 | 2 | 1.01: 6.443; 1.06: 6.8; 1.07: 6.9; 1.08: 5.3; 1.09: 6.8; 1.10: 5.996; 1.11: 6.76; 1.12: 6.3; 1.13: 7.4072; 1.14: 5.7969 |  |  |  | 10
| 1 | 3 | 1.01: 6.3191; 1.06: 6.6; 1.07: 7; 1.08: 6.8; 1.09: 6.9; 1.10: 6.166; 1.11: 6.83; 1.12: 6.3; 1.13: 7.1722; 1.14: 6.0733 |  |  |  | 16
| 1 | 4 | 1.01: 6.3187; 1.06: 6.9; 1.07: 7.1; 1.08: 5.7; 1.09: 6.8; 1.10: 6.042; 1.11: 6.58; 1.12: 6.4; 1.13: 7.5764; 1.14: 6.187 |  |  |  | 6.5436
| 1 | 5 | 1.01: 6.2535; 1.06: 6.8; 1.07: 7; 1.08: 6.2; 1.09: 6.7; 1.10: 6.263; 1.11: 6.48; 1.12: 6.2; 1.13: 7.3132; 1.14: 5.8964 |  |  |  | 56
| 1 | 6 | 1.01: 6.0887; 1.06: 6.95; 1.07: 7.1; 1.08: 6.1; 1.09: 6.7; 1.10: 6.197; 1.11: 6.67; 1.12: 6.1; 1.13: 7.2944; 1.14: 6.0304
| Mean |  | 1.01: 6.2868; 1.06: 6.8083; 1.07: 7.0333; 1.08: 6.05; 1.09: 6.85; 1.10: 6.1378; 1.11: 6.6783; 1.12: 6.25; 1.13: 7.3759; 1.14: 6.0357
| Median |  | 1.01: 6.3082; 1.06: 6.8; 1.07: 7.05; 1.08: 6.15; 1.09: 6.8; 1.10: 6.1645; 1.11: 6.71; 1.12: 6.25; 1.13: 7.3602; 1.14: 6.0519
| Std Dev. |  | 1.01: 0.11568; 1.06: 0.12007; 1.07: 0.08165; 1.08: 0.50892; 1.09: 0.18708; 1.10: 0.099895; 1.11: 0.12952; 1.12: 0.10488; 1.13: 0.14607; 1.14: 0.16641
| Rel.Std.Dev. |  | 1.01: 1.84%; 1.06: 1.764%; 1.07: 1.161%; 1.08: 8.412%; 1.09: 2.731%; 1.10: 1.628%; 1.11: 1.939%; 1.12: 1.678%; 1.13: 1.98%; 1.14: 2.757%
| PDM3 |  | 1.01: -3.925%; 1.06: 4.045%; 1.07: 7.484%; 1.08: -7.543%; 1.09: 4.682%; 1.10: -6.201%; 1.11: 2.059%; 1.12: -4.487%; 1.13: 12.72%; 1.14: -7.762%
| Z-Score (Absolute) |  | 1.01: 0.37; 1.06: 0.62; 1.07: 1.04; 1.08: 0.82; 1.09: 0.7; 1.10: 0.65; 1.11: 0.37; 1.12: 0.44; 1.13: 1.69; 1.14: 0.85
| Analytical results for Sn in OREAS 238b (Certified Value 3.69 ppm) | [27 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4 | 3.6 | 1 | < 10 | 4 | 3.6 | 5.7 | 3.8 | 3.6 | 3.9 | 3.8 | 3.5 | 3.6 | 3.8747 | 3.66 | 3.533 | 3.8 | < 5 | 3.8 | 3.8 | 3.4 | 2.8548 | < 5 | 3.1651 | 3.697 |  | 1
| 1 | 2 | 3.9 | 3 | 1 | < 10 | 3.8 | 3.8 | 3.7 | 4.5 | 3.5 | 4 | 3.9 | 3.6 | 3.6 | 3.8306 | 3.54 | 3.472 | 3.9 | < 5 | 3.7 | 3.4 | 3.4 | 2.8767 | < 5 | 3.2184 | 3.818 |  | 26
| 1 | 3 | 3.9 | 3 | 1 | < 10 | 4 | 3.9 | 4.1 | 4 | 3.6 | 4 | 3.7 | 3.6 | 4 | 3.8142 | 3.57 | 3.636 | 3.8 | < 5 | 3.9 | 3.4 | 3.4 | 2.8739 | < 5 | 3.1862 | 3.7496 |  | 16
| 1 | 4 | 4 | 3.5 | 1 | < 10 | 4 | 4.1 | 4.4 | 3.9 | 3.6 | 3.9 | 4 | 3.5 | 3.7 | 3.7331 | 3.54 | 3.597 | 3.8 | < 5 | 3.8 | 3.3 | 3.4 | 2.8735 | < 5 | 3.2828 | 3.6287 |  | 3.6899
| 1 | 5 | 3.8 | 3.7 | 1 | < 10 | 3.9 | 4.4 | 3.1 | 3.8 | 3.6 | 3.8 | 3.9 | 3.6 | 3.7 | 3.9277 | 3.51 | 3.832 | 3.9 | < 5 | 3.8 | 3.7 | 3.4 | 2.843 | < 5 | 3.2128 | 3.5918 |  | 57
| 1 | 6 | 3.9 | 3.8 | 1 | < 10 | 3.8 | 4.1 | 3 | 3.9 | 3.6 | 3.6 | 4.1 | 3.6 | 3.8 | 3.7615 | 3.5 | 3.558 | 3.8 | < 5 | 3.7 | 3.6 | 3.4 | 2.7945 | < 5 | 3.2872 | 3.5197
| Mean |  | 3.9167 | 3.4333 | 1 |  | 3.9167 | 3.9833 | 4 | 3.9833 | 3.5833 | 3.8667 | 3.9 | 3.5667 | 3.7333 | 3.8236 | 3.5533 | 3.6047 | 3.8333 |  | 3.7833 | 3.5333 | 3.4 | 2.8527 |  | 3.2254 | 3.6675
| Median |  | 3.9 | 3.55 | 1 |  | 3.95 | 4 | 3.9 | 3.9 | 3.6 | 3.9 | 3.9 | 3.6 | 3.7 | 3.8224 | 3.54 | 3.5775 | 3.8 |  | 3.8 | 3.5 | 3.4 | 2.8641 |  | 3.2156 | 3.6629
| Std Dev. |  | 0.075277 | 0.35024 | 0 |  | 0.098319 | 0.27869 | 0.99599 | 0.26394 | 0.040825 | 0.15055 | 0.14142 | 0.05164 | 0.15055 | 0.071593 | 0.05785 | 0.12461 | 0.05164 |  | 0.075277 | 0.19664 | 0 | 0.031445 |  | 0.049989 | 0.10894
| Rel.Std.Dev. |  | 1.922% | 10.2% | 0% |  | 2.51% | 6.996% | 24.9% | 6.626% | 1.139% | 3.894% | 3.626% | 1.448% | 4.033% | 1.872% | 1.628% | 3.457% | 1.347% |  | 1.99% | 5.565% | 0% | 1.102% |  | 1.55% | 2.97%
| PDM3 |  | 6.146% | -6.953% | -72.9% |  | 6.146% | 7.953% | 8.404% | 7.953% | -2.888% | 4.791% | 5.694% | -3.339% | 1.177% | 3.625% | -3.701% | -2.31% | 3.888% |  | 2.533% | -4.243% | -7.856% | -22.69% |  | -12.59% | -0.6077%
| Z-Score (Absolute) |  | 0.7 | 0.87 | NA | 4.21 | 0.7 | 0.92 | 0.97 | 0.92 | 0.38 | 0.54 | 0.65 | 0.43 | 0.11 | 0.4 | 0.48 | 0.31 | 0.43 | 3.89 | 0.27 | 0.54 | 0.97 | 2.75 | 3.89 | 1.54 | 0.11
Results from laboratory 1.06 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 238b (Certified Value 118 ppm) | [27 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 118.41 | 123.4898 | 122 | 121 | 118.4245 | 120.4 | 124 | 119 | 115 | 118.5 | 119.5 | 123.5 | 117.5 | 123 | 113.1971 | 120 | 109.3 | 115 | 115.6 | 117.8 | 118.6 | 113 | 135.3 | 130 | 118.365 | 110.7042 | 1
| 1 | 2 | 118.16 | 124.3816 | 117 | 124 | 120.9743 | 119.9 | 115 | 120 | 115 | 119.5 | 121.5 | 119 | 116 | 116 | 115.3702 | 122 | 108.8 | 118 | 112.8 | 117.2 | 119.3 | 114 | 131.7 | 131 | 119.7782 | 112.9019 | 27
| 1 | 3 | 120.59 | 124.2836 | 126 | 125 | 117.7922 | 120.7 | 118 | 114 | 116 | 119 | 125 | 119.5 | 115.5 | 114 | 116.5443 | 138 | 112.4 | 113 | 114.2 | 119.9 | 118.3 | 114 | 135.6 | 130 | 121.797 | 112.1383 | 16
| 1 | 4 | 122.82 | 123.2546 | 125 | 125 | 119.5397 | 119.9 | 123 | 122 | 118 | 122 | 119 | 122 | 117.5 | 119 | 115.4539 | 119 | 110.9 | 117 | 114.9 | 119 | 118.6 | 115 | 133.9 | 136 | 115.4882 | 111.4667 | 117.9561
| 1 | 5 | 121.07 | 122.9704 | 122 | 121 | 118.6835 | 120.4 | 121 | 111 | 114 | 119.5 | 119 | 117.5 | 120 | 114 | 116.2029 | 116 | 115.4 | 113 | 115.2 | 116.3 | 118.7 | 115 | 134.8 | 129 | 116.3967 | 111.069 | 58
| 1 | 6 | 122.41 | 121.5886 | 122 | 124 | 119.3661 | 120.3 | 116 | 116 | 113 | 123 | 121.5 | 122.5 | 120 | 112 | 113.1847 | 120 | 109.6 | 111 | 112.2 | 115.3 | 119.1 | 113 | 130.7 | 126 | 116.3462 | 113.1696
| Mean |  | 120.5767 | 123.3281 | 122.3333 | 123.3333 | 119.13 | 120.2667 | 119.5 | 117 | 115.1667 | 120.25 | 120.9167 | 120.6667 | 117.75 | 116.3333 | 114.9922 | 122.5 | 111.0667 | 114.5 | 114.15 | 117.5833 | 118.7667 | 114 | 133.6667 | 130.3333 | 118.0285 | 111.9083
| Median |  | 120.83 | 123.3722 | 122 | 124 | 119.0248 | 120.35 | 119.5 | 117.5 | 115 | 119.5 | 120.5 | 120.75 | 117.5 | 115 | 115.412 | 120 | 110.25 | 114 | 114.55 | 117.5 | 118.65 | 114 | 134.35 | 130 | 117.3808 | 111.8025
| Std Dev. |  | 1.9582 | 1.0206 | 3.1411 | 1.8619 | 1.1056 | 0.31411 | 3.7283 | 4.0988 | 1.7224 | 1.8097 | 2.3112 | 2.3381 | 1.917 | 4.0332 | 1.4643 | 7.8422 | 2.4897 | 2.6646 | 1.3708 | 1.6987 | 0.36697 | 0.89443 | 2.0206 | 3.266 | 2.4217 | 0.99803
| Rel.Std.Dev. |  | 1.624% | 0.8276% | 2.568% | 1.51% | 0.9281% | 0.2612% | 3.12% | 3.503% | 1.496% | 1.505% | 1.911% | 1.938% | 1.628% | 3.467% | 1.273% | 6.402% | 2.242% | 2.327% | 1.201% | 1.445% | 0.309% | 0.7846% | 1.512% | 2.506% | 2.052% | 0.8918%
| PDM3 |  | 2.222% | 4.554% | 3.711% | 4.559% | 0.9952% | 1.959% | 1.309% | -0.8106% | -2.365% | 1.945% | 2.51% | 2.298% | -0.1748% | -1.376% | -2.513% | 3.852% | -5.841% | -2.93% | -3.227% | -0.316% | 0.6871% | -3.354% | 13.32% | 10.49% | 0.06139% | -5.127%
| Z-Score (Absolute) |  | 0.37 | 0.98 | 0.76 | 0.99 | 0.04 | 0.3 | 0.12 | 0.44 | 0.85 | 0.29 | 0.44 | 0.39 | 0.27 | 0.59 | 0.89 | 0.8 | 1.77 | 1 | 1.08 | 0.31 | 0.04 | 1.11 | 3.31 | 2.56 | 0.21 | 1.58
| Analytical results for Ta in OREAS 238b (Certified Value 1.07 ppm) | [27 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | [5 empty] | 2
3
| 1 | 1 | 1.18 | 1.0176 | < 0.1 | < 5 | 1.05 | 0.99 | 0.85 | 0.7 | 1.07 | 1.1 | 1.05 | 1.07 | 1 | 1.1358 | 1.02 | 0.94 | 1.2 | 1.14 | 0.8 | 0.6181 | 1.2094 | [5 empty] | 1
| 1 | 2 | 1.14 | 0.96559 | < 0.1 | 5.05 | 1.03 | 0.92 | 0.94 | 0.7 | 1.08 | 1.14 | 1.1 | 1.04 | 0.9 | 1.1552 | 0.99 | 1.03 | 1.22 | 1.04 | 0.9 | 0.6861 | 1.1687 | [5 empty] | 28
| 1 | 3 | 1.16 | 0.97048 | < 0.1 | < 5 | 1.04 | 1.08 | 0.9 | 0.7 | 1.06 | 1.15 | 1.05 | 1.04 | 1.1 | 1.2242 | 1.01 | 0.99 | 1.24 | 1.04 | 0.9 | 0.6048 | 1.173 | [5 empty] | 16
| 1 | 4 | 1.15 | 1.0532 | < 0.1 | < 5 | 1.08 | 1.06 | 0.99 | 0.8 | 1.1 | 1.1 | 1.03 | 1.08 | 1 | 1.2259 | 0.98 | 0.96 | 1.2 | 1.02 | 0.9 | 0.6221 | 1.1556 | [5 empty] | 1.0702
| 1 | 5 | 1.08 | 1.024 | < 0.1 | 5.0767 | 1.08 | 1.11 | 0.87 | 0.6 | 1.11 | 1.08 | 1.04 | 1.08 | 1.1 | 1.2416 | 0.98 | 0.98 | 1.17 | 1.11 | 0.8 | 0.6229 | 1.1599 | [5 empty] | 59
| 1 | 6 | 1.14 | 0.99774 | 0.1 | < 5 | 1.06 | 0.95 | 0.95 | 0.7 | 1.09 | 1.11 | 1.08 | 1.08 | 1.1 | 1.1538 | 0.98 | 1 | 1.14 | 1.07 | 0.8 | 0.6259 | 1.1543
| Mean |  | 1.1417 | 1.0048 | 0.1 | 5.0633 | 1.0567 | 1.0183 | 0.91667 | 0.7 | 1.085 | 1.1133 | 1.0583 | 1.065 | 1.0333 | 1.1894 | 0.99333 | 0.98333 | 1.195 | 1.07 | 0.85 | 0.62998 | 1.1701
| Median |  | 1.145 | 1.0077 | 0.1 | 5.0633 | 1.055 | 1.025 | 0.92 | 0.7 | 1.085 | 1.105 | 1.05 | 1.075 | 1.05 | 1.1897 | 0.985 | 0.985 | 1.2 | 1.055 | 0.85 | 0.6225 | 1.1643
| Std Dev. |  | 0.033714 | 0.033575 |  | 0.018856 | 0.020656 | 0.076267 | 0.052789 | 0.063246 | 0.018708 | 0.026583 | 0.026394 | 0.019748 | 0.08165 | 0.045988 | 0.017512 | 0.031411 | 0.035637 | 0.046476 | 0.054772 | 0.028473 | 0.020572
| Rel.Std.Dev. |  | 2.953% | 3.342% |  | 0.3724% | 1.955% | 7.489% | 5.759% | 9.035% | 1.724% | 2.388% | 2.494% | 1.854% | 7.902% | 3.866% | 1.763% | 3.194% | 2.982% | 4.344% | 6.444% | 4.52% | 1.758%
| PDM3 |  | 6.677% | -6.116% | -90.66% | 373.1% | -1.265% | -4.847% | -14.35% | -34.59% | 1.382% | 4.029% | -1.11% | -0.4868% | -3.446% | 11.14% | -7.183% | -8.118% | 11.66% | -0.0196% | -20.58% | -41.13% | 9.337%
| Z-Score (Absolute) |  | 0.72 | 0.57 | NA | 21.54 | 0.08 | 0.44 | 1.4 | NA | 0.19 | 0.45 | 0.06 | 0 | NA | 1.17 | 0.67 | 0.77 | 1.22 | 0.05 | NA | 4.09 | 0.99
Results from laboratories 1.06, 1.11, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 238b (Certified Value 0.67 ppm) | [27 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.21; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.34 | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*OES; 1.13: 4A*MS | 2
3
| 1 | 1 | 1.01: 0.67949; 1.03: 0.68; 1.05: 0.7; 1.06: 0.7; 1.07: 0.71; 1.08: 0.652; 1.09: 0.61; 1.10: 0.72; 1.11: 0.6; 1.12: 2.23; 1.13: 5.2051 | 1
| 1 | 2 | 1.01: 0.68684; 1.03: 0.68; 1.05: 0.8; 1.06: 0.5; 1.07: 0.68; 1.08: 0.633; 1.09: 0.64; 1.10: 0.74; 1.11: 0.6; 1.12: 2.34; 1.13: 4.8195 | 11
| 1 | 3 | 1.01: 0.6651; 1.03: 0.66; 1.05: 0.8; 1.06: 0.7; 1.07: 0.69; 1.08: 0.657; 1.09: 0.62; 1.10: 0.71; 1.11: 0.6; 1.12: 2.09; 1.13: 5.0996 | 16
| 1 | 4 | 1.01: 0.65054; 1.03: 0.72; 1.05: 0.8; 1.06: 0.6; 1.07: 0.67; 1.08: 0.676; 1.09: 0.61; 1.10: 0.72; 1.11: 0.6; 1.12: 2; 1.13: 4.9319 | 0.67372
| 1 | 5 | 1.01: 0.67743; 1.03: 0.7; 1.05: 0.7; 1.06: 0.6; 1.07: 0.68; 1.08: 0.651; 1.09: 0.65; 1.10: 0.7; 1.11: 0.6; 1.12: 1.9; 1.13: 4.9111 | 60
| 1 | 6 | 1.01: 0.67864; 1.03: 0.72; 1.05: 0.8; 1.06: 0.7; 1.07: 0.68; 1.08: 0.657; 1.09: 0.62; 1.10: 0.68; 1.11: 0.5; 1.12: 1.98; 1.13: 4.8875
| Mean |  | 1.01: 0.67301; 1.03: 0.69333; 1.05: 0.76667; 1.06: 0.63333; 1.07: 0.685; 1.08: 0.65433; 1.09: 0.625; 1.10: 0.71167; 1.11: 0.58333; 1.12: 2.09; 1.13: 4.9758
| Median |  | 1.01: 0.67804; 1.03: 0.69; 1.05: 0.8; 1.06: 0.65; 1.07: 0.68; 1.08: 0.6545; 1.09: 0.62; 1.10: 0.715; 1.11: 0.6; 1.12: 2.045; 1.13: 4.9215
| Std Dev. |  | 1.01: 0.013056; 1.03: 0.024221; 1.05: 0.05164; 1.06: 0.08165; 1.07: 0.013784; 1.08: 0.013823; 1.09: 0.016432; 1.10: 0.020412; 1.11: 0.040825; 1.12: 0.16637; 1.13: 0.14576
| Rel.Std.Dev. |  | 1.01: 1.94%; 1.03: 3.493%; 1.05: 6.736%; 1.06: 12.89%; 1.07: 2.012%; 1.08: 2.112%; 1.09: 2.629%; 1.10: 2.868%; 1.11: 6.999%; 1.12: 7.96%; 1.13: 2.929%
| PDM3 |  | 1.01: -0.1063%; 1.03: 2.911%; 1.05: 13.8%; 1.06: -5.995%; 1.07: 1.674%; 1.08: -2.878%; 1.09: -7.232%; 1.10: 5.632%; 1.11: -13.42%; 1.12: 210.2%; 1.13: 638.5%
| Z-Score (Absolute) |  | 1.01: 0.38; 1.03: 0.1; 1.05: NA; 1.06: NA; 1.07: 0.1; 1.08: 0.82; 1.09: 1.51; 1.10: 0.53; 1.11: NA; 1.12: 32.95; 1.13: 100.83
Results from laboratories 1.06, 1.16 and 1.28 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 238b (Indicative Value 0.056 ppm) | [27 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 3
3
| 1 | 1 | < 0.2 | < 0.2 | < 0.1 | 10.292 | < 0.1 | 0.09 | < 0.1 | 0.1 | < 0.05 | < 0.05 | 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.07491 | [8 empty] | 1
| 1 | 2 | < 0.2 | < 0.2 | < 0.1 | 10.0387 | < 0.1 | 0.11 | 0.1 | 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.05995 | [8 empty] | 5
| 1 | 3 | < 0.2 | < 0.2 | < 0.1 | 9.5867 | < 0.1 | 0.08 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | 0.06 | < 0.05 | 0.06 | 0.06 | < 0.05 | 0.07491 | [8 empty] | 16
| 1 | 4 | < 0.2 | < 0.2 | < 0.1 | 9.8133 | < 0.1 | 0.09 | 0.2 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | 0.04 | 0.05 | < 0.05 | 0.04499 | [8 empty] | 0.056329
| 1 | 5 | < 0.2 | < 0.2 | < 0.1 | 8.6153 | < 0.1 | 0.08 | 0.3 | < 0.1 | < 0.05 | < 0.05 | 0.06 | 0.05 | 0.07 | < 0.05 | < 0.03 | 0.07 | < 0.05 | 0.01496 | [8 empty] | 11
| 1 | 6 | < 0.2 | < 0.2 | < 0.1 | 8.8473 | < 0.1 | 0.11 | < 0.1 | < 0.1 | < 0.05 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.04 | 0.07 | < 0.05 | 0.01507
| Mean |  | 1.07: 9.5322; 1.09: 0.093333; 1.10: 0.2; 1.11: 0.1; 1.14: 0.055; 1.15: 0.051667; 1.16: 0.065; 1.26: 0.046667; 1.27: 0.065; 1.30: 0.047465
| Median |  | 1.07: 9.7; 1.09: 0.09; 1.10: 0.2; 1.11: 0.1; 1.14: 0.055; 1.15: 0.05; 1.16: 0.065; 1.26: 0.04; 1.27: 0.07; 1.30: 0.05247
| Std Dev. |  | 1.07: 0.66716; 1.09: 0.013663; 1.10: 0.1; 1.11: 0; 1.14: 0.0070711; 1.15: 0.0040825; 1.16: 0.0070711; 1.26: 0.011547; 1.27: 0.0083666; 1.30: 0.027475
| Rel.Std.Dev. |  | 1.07: 6.999%; 1.09: 14.64%; 1.10: 50%; 1.11: 0%; 1.14: 12.86%; 1.15: 7.902%; 1.16: 10.88%; 1.26: 24.74%; 1.27: 12.87%; 1.30: 57.89%
| PDM3 |  | 1.07: 16822%; 1.09: 65.69%; 1.10: 255.1%; 1.11: 77.53%; 1.14: -2.36%; 1.15: -8.278%; 1.16: 15.39%; 1.26: -17.15%; 1.27: 15.39%; 1.30: -15.74%
| Z-Score (Absolute) |  | 1.53 | 1.53 | 0 | 289.54 | 0 | 1.32 | 2.29 | 0.51 | 0.76 | 0.76 | 0.46 | 0.05 | 0.36 | 0.76 | 0.59 | 0.46 | 0.76 | 0.08
| Analytical results for Th in OREAS 238b (Certified Value 14.5 ppm) | [27 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 1
2
| 1 | 1 | 14.76 | 15.0688 | 18.4 | 15.2 | 14.96 | 12.47 | 14.9 | 14.4 | 15 | 13.9 | 14.3 | 15.25 | 17.6 | 13.9174 | 14 | 13.86 | 14.1 | 14.1 | 16 | 13.7 | 12.8958 | 15.0866 | 13.9082 |  |  |  | 1
| 1 | 2 | 15 | 15.1711 | 18.2 | 15.6 | 14.55 | 13.75 | 15.2 | 14.7 | 15.05 | 15.25 | 14.3 | 14.95 | 13.2 | 14.4688 | 13.4 | 14.2 | 14.1 | 14.1 | 14.7 | 14.2 | 13.2018 | 15.126 | 12.9107 |  |  |  | 31
| 1 | 3 | 15.01 | 15.2787 | 18.3 | 15.3 | 14.77 | 14.52 | 15.2 | 14.5 | 13.95 | 15.45 | 14.6 | 15.25 | 20.4 | 14.5816 | 13.8 | 14.4 | 13.9 | 14.4 | 14.6 | 14.2 | 13.2132 | 15.4019 | 13.5955 |  |  |  | 16
| 1 | 4 | 15.3 | 15.4339 | 17.8 | 15.1 | 14.67 | 15.34 | 15.6 | 14.5 | 14.55 | 15.95 | 14.4 | 15.1 | 15.7 | 14.1424 | 13.4 | 14.13 | 13.9 | 14 | 14 | 14.2 | 13.2258 | 14.259 | 13.502 |  |  |  | 14.513
| 1 | 5 | 15.13 | 15.392 | 18.1 | 15.3 | 14.98 | 14.73 | 14.5 | 14.4 | 14.25 | 15.45 | 13.95 | 15.7 | 17.6 | 14.2385 | 13.5 | 14.17 | 14.2 | 13.8 | 15.5 | 14 | 13.1857 | 14.6334 | 13.4575 |  |  |  | 61
| 1 | 6 | 15.27 | 15.0487 | 18.6 | 15.3 | 14.6 | 14.78 | 15.6 | 14.4 | 15 | 14.75 | 14.8 | 15.4 | 17 | 14.338 | 13.5 | 14.18 | 14.2 | 13.2 | 15.3 | 14 | 12.7623 | 14.7319 | 13.2617
| Mean |  | 15.0783 | 15.2322 | 18.2333 | 15.3 | 14.755 | 14.265 | 15.1667 | 14.4833 | 14.6333 | 15.125 | 14.3917 | 15.275 | 16.9167 | 14.2811 | 13.6 | 14.1567 | 14.0667 | 13.9333 | 15.0167 | 14.05 | 13.0808 | 14.8731 | 13.4393
| Median |  | 15.07 | 15.2249 | 18.25 | 15.3 | 14.72 | 14.625 | 15.2 | 14.45 | 14.775 | 15.35 | 14.35 | 15.25 | 17.3 | 14.2882 | 13.5 | 14.175 | 14.1 | 14.05 | 15 | 14.1 | 13.1937 | 14.9093 | 13.4797
| Std Dev. |  | 0.20034 | 0.16279 | 0.27325 | 0.16733 | 0.18229 | 1.0188 | 0.42269 | 0.1169 | 0.46116 | 0.71397 | 0.29055 | 0.25836 | 2.3819 | 0.23753 | 0.24495 | 0.17351 | 0.13663 | 0.40825 | 0.71949 | 0.19748 | 0.19992 | 0.41078 | 0.33451
| Rel.Std.Dev. |  | 1.329% | 1.069% | 1.499% | 1.094% | 1.235% | 7.142% | 2.787% | 0.8072% | 3.151% | 4.72% | 2.019% | 1.691% | 14.08% | 1.663% | 1.801% | 1.226% | 0.9713% | 2.93% | 4.791% | 1.406% | 1.528% | 2.762% | 2.489%
| PDM3 |  | 3.896% | 4.956% | 25.63% | 5.423% | 1.668% | -1.709% | 4.504% | -0.2043% | 0.8293% | 4.217% | -0.8359% | 5.251% | 16.56% | -1.598% | -6.291% | -2.455% | -3.075% | -3.994% | 3.471% | -3.19% | -9.868% | 2.482% | -7.398%
| Z-Score (Absolute) |  | 0.56 | 0.76 | 4.55 | 0.84 | 0.15 | 0.47 | 0.67 | 0.19 | 0 | 0.62 | 0.31 | 0.81 | 2.89 | 0.45 | 1.31 | 0.6 | 0.72 | 0.89 | 0.48 | 0.74 | 1.96 | 0.3 | 1.51
| Analytical results for Ti in OREAS 238b (Certified Value 0.461 wt.%) | [27 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | wt.%
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.4763 | 0.44493 | 0.46 | 0.432 | 0.35 | 0.502 | 0.45 | 0.442 | 0.464 | 0.459 | 0.461 | 0.462 | 0.456 | 0.4794 | 0.46011 | 0.5245 | 0.46 | 0.45 | 0.46 | 0.5 | 0.436 | 0.37024 | 0.5048 | 0.4354 | 0.4602 |  | 1
| 1 | 2 | 0.4709 | 0.45913 | 0.475 | 0.443 | 0.34 | 0.508 | 0.45 | 0.445 | 0.467 | 0.466 | 0.474 | 0.444 | 0.458 | 0.4651 | 0.47564 | 0.5365 | 0.47 | 0.44 | 0.45 | 0.5 | 0.442 | 0.36759 | 0.5027 | 0.4478 | 0.48018 |  | 32
| 1 | 3 | 0.4766 | 0.4543 | 0.47 | 0.423 | 0.34 | 0.514 | 0.46 | 0.44 | 0.455 | 0.459 | 0.475 | 0.454 | 0.452 | 0.4712 | 0.46477 | 0.6055 | 0.45 | 0.45 | 0.46 | 0.51 | 0.441 | 0.38103 | 0.5043 | 0.4565 | 0.46724 |  | 16
| 1 | 4 | 0.4767 | 0.4586 | 0.465 | 0.425 | 0.34667 | 0.516 | 0.45 | 0.445 | 0.455 | 0.472 | 0.458 | 0.461 | 0.46 | 0.4774 | 0.47036 | 0.5233 | 0.46 | 0.45 | 0.46 | 0.5 | 0.444 | 0.3754 | 0.5268 | 0.3869 | 0.47329 |  | 0.46128
| 1 | 5 | 0.4557 | 0.45065 | 0.47 | 0.435 | 0.35333 | 0.514 | 0.45 | 0.424 | 0.449 | 0.46 | 0.451 | 0.452 | 0.474 | 0.4733 | 0.47553 | 0.512 | 0.44 | 0.45 | 0.43 | 0.5 | 0.452 | 0.37283 | 0.5001 | 0.3927 | 0.46965 |  | 62
| 1 | 6 | 0.4736 | 0.44927 | 0.465 | 0.447 | 0.35333 | 0.502 | 0.45 | 0.44 | 0.453 | 0.48 | 0.484 | 0.465 | 0.467 | 0.457 | 0.46629 | 0.5233 | 0.44 | 0.45 | 0.45 | 0.5 | 0.442 | 0.36245 | 0.4917 | 0.3888 | 0.47645
| Mean |  | 0.47163 | 0.45281 | 0.4675 | 0.43417 | 0.34722 | 0.50933 | 0.45167 | 0.43933 | 0.45717 | 0.466 | 0.46717 | 0.45633 | 0.46117 | 0.47057 | 0.46878 | 0.53752 | 0.45333 | 0.44833 | 0.45167 | 0.50167 | 0.44283 | 0.37159 | 0.50507 | 0.41802 | 0.47117
| Median |  | 0.47495 | 0.45248 | 0.4675 | 0.4335 | 0.34833 | 0.511 | 0.45 | 0.441 | 0.455 | 0.463 | 0.4675 | 0.4575 | 0.459 | 0.47225 | 0.46832 | 0.5239 | 0.455 | 0.45 | 0.455 | 0.5 | 0.442 | 0.37154 | 0.5035 | 0.41405 | 0.47147
| Std Dev. |  | 0.0081289 | 0.0055674 | 0.005244 | 0.0095586 | 0.0061162 | 0.0062823 | 0.0040825 | 0.0078401 | 0.0068823 | 0.0085557 | 0.012449 | 0.0078145 | 0.0080104 | 0.0083251 | 0.0062065 | 0.034198 | 0.012111 | 0.0040825 | 0.01169 | 0.0040825 | 0.0052313 | 0.0064269 | 0.011678 | 0.032041 | 0.0070946
| Rel.Std.Dev. |  | 1.724% | 1.23% | 1.122% | 2.202% | 1.761% | 1.233% | 0.9039% | 1.785% | 1.505% | 1.836% | 2.665% | 1.712% | 1.737% | 1.769% | 1.324% | 6.362% | 2.671% | 0.9106% | 2.588% | 0.8138% | 1.181% | 1.73% | 2.312% | 7.665% | 1.506%
| PDM3 |  | 2.245% | -1.835% | 1.349% | -5.877% | -24.73% | 10.42% | -2.084% | -4.757% | -0.8913% | 1.024% | 1.277% | -1.072% | -0.02413% | 2.014% | 1.627% | 16.53% | -1.722% | -2.806% | -2.084% | 8.756% | -3.999% | -19.44% | 9.493% | -9.379% | 2.144%
| Z-Score (Absolute) |  | 0.73 | 0.22 | 0.52 | 1.16 | 5.53 | 2.63 | 0.28 | 0.9 | 0 | 0.44 | 0.5 | 0.04 | 0.2 | 0.67 | 0.58 | 4.04 | 0.19 | 0.44 | 0.28 | 2.24 | 0.72 | 4.31 | 2.41 | 1.97 | 0.7
| Analytical results for Tl in OREAS 238b (Certified Value 0.85 ppm) | [27 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 0.87 | 0.79527 | 0.8 | 0.89 | 0.86 | 0.89 | 0.83 | 0.9 | 0.9 | 0.82 | 0.86 | 0.82 | 0.81 | 0.89735 | 0.798 | 0.79 | 0.85 | 0.84 | 0.79 | 0.94622 | 0.9041 | 0.9037 |  |  |  |  | 1
| 1 | 2 | 0.88 | 0.8195 | 0.8 | 0.91 | 0.84 | 0.95 | 0.79 | 0.9 | 0.9 | 0.83 | 0.83 | 0.82 | 0.77 | 0.86792 | 0.789 | 0.82 | 0.87 | 0.74 | 0.82 | 0.93885 | 0.8937 | 0.85708 |  |  |  |  | 33
| 1 | 3 | 0.89 | 0.82581 | 0.8 | 0.89 | 0.87 | 0.96 | 0.84 | 0.9 | 0.9 | 0.86 | 0.8 | 0.81 | 0.81 | 0.90622 | 0.816 | 0.79 | 0.85 | 0.74 | 0.82 | 0.91971 | 0.8742 | 0.87838 |  |  |  |  | 16
| 1 | 4 | 0.92 | 0.81924 | 0.9 | 0.9 | 0.85 | 0.85 | 0.8 | 0.9 | 0.9 | 0.84 | 0.77 | 0.86 | 0.77 | 0.89598 | 0.802 | 0.78 | 0.86 | 0.71 | 0.82 | 0.92048 | 0.8779 | 0.86062 |  |  |  |  | 0.8481
| 1 | 5 | 0.9 | 0.82227 | 0.8 | 0.89 | 0.86 | 0.84 | 0.78 | 0.9 | 0.94 | 0.83 | 0.82 | 0.85 | 0.85 | 0.87366 | 0.838 | 0.8 | 0.85 | 0.82 | 0.8 | 0.94325 | 0.8486 | 0.85879 |  |  |  |  | 63
| 1 | 6 | 0.89 | 0.78353 | 0.8 | 0.89 | 0.86 | 0.9 | 0.81 | 0.9 | 0.95 | 0.86 | 0.84 | 0.83 | 0.85 | 0.92748 | 0.781 | 0.78 | 0.82 | 0.78 | 0.79 | 0.89628 | 0.8657 | 0.84822
| Mean |  | 0.89167 | 0.81094 | 0.81667 | 0.895 | 0.85667 | 0.89833 | 0.80833 | 0.9 | 0.915 | 0.84 | 0.82 | 0.83167 | 0.81 | 0.89477 | 0.804 | 0.79333 | 0.85 | 0.77167 | 0.80667 | 0.92746 | 0.87737 | 0.8678
| Median |  | 0.89 | 0.81937 | 0.8 | 0.89 | 0.86 | 0.895 | 0.805 | 0.9 | 0.9 | 0.835 | 0.825 | 0.825 | 0.81 | 0.89667 | 0.8 | 0.79 | 0.85 | 0.76 | 0.81 | 0.92966 | 0.87605 | 0.85971
| Std Dev. |  | 0.017224 | 0.017253 | 0.040825 | 0.0083666 | 0.010328 | 0.049565 | 0.023166 | 0 | 0.023452 | 0.016733 | 0.031623 | 0.019408 | 0.035777 | 0.0218 | 0.020465 | 0.015055 | 0.016733 | 0.050761 | 0.015055 | 0.019037 | 0.019776 | 0.020158
| Rel.Std.Dev. |  | 1.932% | 2.128% | 4.999% | 0.9348% | 1.206% | 5.517% | 2.866% | 0% | 2.563% | 1.992% | 3.856% | 2.334% | 4.417% | 2.436% | 2.545% | 1.898% | 1.969% | 6.578% | 1.866% | 2.053% | 2.254% | 2.323%
| PDM3 |  | 5.137% | -4.381% | -3.706% | 5.53% | 1.01% | 5.923% | -4.689% | 6.12% | 7.888% | -0.9549% | -3.313% | -1.938% | -4.492% | 5.503% | -5.2% | -6.457% | 0.2242% | -9.012% | -4.885% | 9.358% | 3.451% | 2.323%
| Z-Score (Absolute) |  | 0.84 | 0.61 | NA | 0.9 | 0.21 | 0.96 | 0.66 | NA | 1.26 | 0.09 | 0.45 | 0.24 | 0.63 | 0.89 | 0.74 | 0.93 | 0.09 | 1.32 | 0.69 | 1.48 | 0.58 | 0.41
Results from laboratories 1.05 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 238b (Certified Value 0.27 ppm) | [27 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.20; 1.08: 1.26; 1.09: 1.28; 1.10: 1.30; 1.11: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*MS; 1.09: 4A*OES/MS; 1.10: 4A*MS; 1.11: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.2388; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.206; 1.08: 0.28; 1.09: 0.2; 1.10: 0.1935; 1.11: 0.40235 |  | 1
| 1 | 2 | 1.01: 0.2325; 1.03: 0.3; 1.05: 0.3; 1.06: 0.2; 1.07: 0.208; 1.08: 0.3; 1.09: 0.2; 1.10: 0.2051; 1.11: 0.37186 |  | 13
| 1 | 3 | 1.01: 0.23217; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.216; 1.08: 0.3; 1.09: 0.2; 1.10: 0.1979; 1.11: 0.38911 |  | 16
| 1 | 4 | 1.01: 0.23512; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.21; 1.08: 0.3; 1.09: 0.2; 1.10: 0.1998; 1.11: 0.37894 |  | 0.26676
| 1 | 5 | 1.01: 0.23392; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.208; 1.08: 0.27; 1.09: 0.2; 1.10: 0.205; 1.11: 0.37612 |  | 64
| 1 | 6 | 1.01: 0.23752; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.216; 1.08: 0.28; 1.09: 0.2; 1.11: 0.38086
| Mean |  | 1.01: 0.235; 1.03: 0.3; 1.05: 0.3; 1.06: 0.28333; 1.07: 0.21067; 1.08: 0.28833; 1.09: 0.2; 1.10: 0.20026; 1.11: 0.38321
| Median |  | 1.01: 0.23452; 1.03: 0.3; 1.05: 0.3; 1.06: 0.3; 1.07: 0.209; 1.08: 0.29; 1.09: 0.2; 1.10: 0.1998; 1.11: 0.3799
| Std Dev. |  | 1.01: 0.0026911; 1.03: 0; 1.05: 0; 1.06: 0.040825; 1.07: 0.0043205; 1.08: 0.013292; 1.09: 3.0405e-17; 1.10: 0.0049339; 1.11: 0.010991
| Rel.Std.Dev. |  | 1.01: 1.145%; 1.03: 0%; 1.05: 0%; 1.06: 14.41%; 1.07: 2.051%; 1.08: 4.61%; 1.09: 1.52e-14%; 1.10: 2.464%; 1.11: 2.868%
| PDM3 |  | 1.01: -11.9%; 1.03: 12.46%; 1.05: 12.46%; 1.06: 6.214%; 1.07: -21.03%; 1.08: 8.089%; 1.09: -25.03%; 1.10: -24.93%; 1.11: 43.65%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.03: 0.23; 1.05: 0.23; 1.06: 0; 1.07: 1.01; 1.08: 0.07; 1.09: 1.16; 1.10: 1.16; 1.11: 1.39
| Analytical results for U in OREAS 238b (Certified Value 2.74 ppm) | [27 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
3
| 1 | 1 | 2.73 | 2.7726 | 2.8 | 2.9 | 2.82 | 2.6 | 2.99 | 2.8 | 2.6 | 2.8 | 2.8 | 2.6 | 2.6 | 2.717 | 2.69 | 2.514 | 2.6 | 2.74 | 2.9 | 2.5 | 2.498 | 3.9777 | 3.0602 |  |  |  | 1
| 1 | 2 | 2.77 | 2.7717 | 2.9 | 2.9 | 2.75 | 2.8 | 3 | 2.8 | 2.6 | 2.8 | 2.7 | 2.6 | 2.5 | 2.6857 | 2.7 | 2.54 | 2.74 | 2.82 | 2.6 | 2.5 | 2.5875 | 4.1841 | 2.897 |  |  |  | 35
| 1 | 3 | 2.78 | 2.8395 | 3 | 2.9 | 2.83 | 2.8 | 3.08 | 2.8 | 2.6 | 2.8 | 2.8 | 2.6 | 3.1 | 2.6907 | 2.64 | 2.596 | 2.64 | 2.83 | 2.5 | 2.6 | 2.5598 | 4.014 | 3.0894 |  |  |  | 16
| 1 | 4 | 2.75 | 2.7927 | 2.9 | 2.8 | 2.78 | 2.9 | 3.06 | 2.7 | 2.7 | 2.8 | 2.7 | 2.6 | 2.6 | 2.7347 | 2.65 | 2.6 | 2.62 | 2.78 | 2.4 | 2.6 | 2.5534 | 4.1053 | 3.0174 |  |  |  | 2.7363
| 1 | 5 | 2.78 | 2.6789 | 2.9 | 2.8 | 2.82 | 2.8 | 2.85 | 2.7 | 2.7 | 2.7 | 2.6 | 2.7 | 2.8 | 2.778 | 2.63 | 2.656 | 2.62 | 2.75 | 2.7 | 2.5 | 2.5847 | 4.1573 | 3.0278 |  |  |  | 65
| 1 | 6 | 2.77 | 2.7018 | 2.9 | 2.9 | 2.75 | 2.9 | 2.93 | 2.7 | 2.7 | 2.7 | 2.9 | 2.6 | 2.7 | 2.7137 | 2.62 | 2.627 | 2.66 | 2.68 | 2.7 | 2.5 | 2.5044 | 4.0439 | 2.9575
| Mean |  | 2.7633 | 2.7595 | 2.9 | 2.8667 | 2.7917 | 2.8 | 2.985 | 2.75 | 2.65 | 2.7667 | 2.75 | 2.6167 | 2.7167 | 2.72 | 2.655 | 2.5888 | 2.6467 | 2.7667 | 2.6333 | 2.5333 | 2.548 | 4.0804 | 3.0082
| Median |  | 2.77 | 2.7721 | 2.9 | 2.9 | 2.8 | 2.8 | 2.995 | 2.75 | 2.65 | 2.8 | 2.75 | 2.6 | 2.65 | 2.7154 | 2.645 | 2.598 | 2.63 | 2.765 | 2.65 | 2.5 | 2.5566 | 4.0746 | 3.0226
| Std Dev. |  | 0.019664 | 0.05943 | 0.063246 | 0.05164 | 0.03656 | 0.10954 | 0.085029 | 0.054772 | 0.054772 | 0.05164 | 0.10488 | 0.040825 | 0.2137 | 0.033632 | 0.032711 | 0.053158 | 0.050067 | 0.055737 | 0.17512 | 0.05164 | 0.038669 | 0.081946 | 0.070281
| Rel.Std.Dev. |  | 0.7116% | 2.154% | 2.181% | 1.801% | 1.31% | 3.912% | 2.849% | 1.992% | 2.067% | 1.866% | 3.814% | 1.56% | 7.866% | 1.236% | 1.232% | 2.053% | 1.892% | 2.015% | 6.65% | 2.038% | 1.518% | 2.008% | 2.336%
| PDM3 |  | 0.9898% | 0.8513% | 5.984% | 4.766% | 2.025% | 2.33% | 9.091% | 0.5025% | -3.152% | 1.112% | 0.5025% | -4.37% | -0.7157% | -0.5947% | -2.969% | -5.388% | -3.274% | 1.112% | -3.761% | -7.416% | -6.881% | 49.12% | 9.938%
| Z-Score (Absolute) |  | 0.09 | 0.06 | 1.01 | 0.79 | 0.28 | 0.34 | 1.58 | 0 | 0.67 | 0.11 | 0 | 0.9 | 0.22 | 0.2 | 0.64 | 1.09 | 0.7 | 0.11 | 0.79 | 1.46 | 1.36 | 8.97 | 1.74
| Analytical results for V in OREAS 238b (Certified Value 108 ppm) | [27 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 109 | 91.12 | 105 | 118 | 118.2393 | 109 | 107 | 100 | 105 | 107 | 110 | 107 | 104 | 111 | 109.6649 | 108 | 110 | 100.6 | 109 | 109 | 106 | 100.21 | 112 | 108.8493 | 102.8425 |  | 1
| 1 | 2 | 106 | 92.12 | 105 | 120 | 118.7589 | 110 | 107 | 107 | 106 | 109 | 113 | 104 | 103 | 108 | 108.3296 | 110 | 111 | 101.2 | 108 | 110 | 108 | 97.68 | 112 | 109.3319 | 105.1227 |  | 36
| 1 | 3 | 111 | 91.07 | 105 | 116 | 118.3204 | 113 | 107 | 105 | 104 | 107 | 114 | 104 | 102 | 110 | 107.7522 | 120 | 111 | 100.9 | 109 | 111 | 106 | 100.21 | 112 | 108.9297 | 104.1315 |  | 16
| 1 | 4 | 112 | 90.56 | 110 | 115 | 117.7688 | 111 | 107 | 104 | 108 | 110 | 110 | 106 | 107 | 111 | 108.247 | 106 | 111 | 102.3 | 109 | 110 | 108 | 99.66 | 117 | 75.3883 | 106.2577 |  | 107.6332
| 1 | 5 | 111 | 90.36 | 105 | 120 | 120.8128 | 110 | 107 | 103 | 105 | 108 | 109 | 106 | 108 | 109 | 107.6871 | 103 | 110 | 101.3 | 106 | 110 | 107 | 99.44 | 111 | 107.2406 | 106.5306 |  | 66
| 1 | 6 | 110 | 92.1 | 110 | 122 | 121.0391 | 112 | 106 | 102 | 110 | 111 | 111 | 107 | 106 | 109 | 106.917 | 108 | 109 | 101.1 | 108 | 110 | 107 | 97.79 | 109 | 105.2297 | 106.2008
| Mean |  | 109.8333 | 91.2217 | 106.6667 | 118.5 | 119.1566 | 110.8333 | 106.8333 | 103.5 | 106.3333 | 108.6667 | 111.1667 | 105.6667 | 105 | 109.6667 | 108.0996 | 109.1667 | 110.3333 | 101.2333 | 108.1667 | 110 | 107 | 99.165 | 112.1667 | 102.4949 | 105.181
| Median |  | 110.5 | 91.095 | 105 | 119 | 118.5397 | 110.5 | 107 | 103.5 | 105.5 | 108.5 | 110.5 | 106 | 105 | 109.5 | 107.9996 | 108 | 110.5 | 101.15 | 108.5 | 110 | 107 | 99.55 | 112 | 108.045 | 105.6617
| Std Dev. |  | 2.137 | 0.74727 | 2.582 | 2.6646 | 1.408 | 1.472 | 0.40825 | 2.429 | 2.2509 | 1.633 | 1.9408 | 1.3663 | 2.3664 | 1.2111 | 0.91806 | 5.8109 | 0.8165 | 0.5785 | 1.169 | 0.63246 | 0.89443 | 1.149 | 2.6394 | 13.3664 | 1.4556
| Rel.Std.Dev. |  | 1.946% | 0.8192% | 2.421% | 2.249% | 1.182% | 1.328% | 0.3821% | 2.347% | 2.117% | 1.503% | 1.746% | 1.293% | 2.254% | 1.104% | 0.8493% | 5.323% | 0.74% | 0.5715% | 1.081% | 0.575% | 0.8359% | 1.159% | 2.353% | 13.04% | 1.384%
| PDM3 |  | 2.044% | -15.25% | -0.898% | 10.1% | 10.71% | 2.973% | -0.7431% | -3.84% | -1.208% | 0.9602% | 3.283% | -1.827% | -2.446% | 1.889% | 0.4334% | 1.425% | 2.509% | -5.946% | 0.4956% | 2.199% | -0.5883% | -7.868% | 4.212% | -4.774% | -2.278%
| Z-Score (Absolute) |  | 0.48 | 4.68 | 0.4 | 2.88 | 3.06 | 0.76 | 0.35 | 1.27 | 0.49 | 0.16 | 0.85 | 0.67 | 0.86 | 0.43 | 0 | 0.3 | 0.62 | 1.9 | 0.02 | 0.53 | 0.3 | 2.48 | 1.13 | 1.55 | 0.81
| Analytical results for W in OREAS 238b (Certified Value 2.8 ppm) | [27 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 3.3 | 2.9126 | 2.5 | < 0.1 | 19.838 | 2.8 | 2.2 | 8.8 | 2.4 | 3 | 2.8 | 2.7 | 2.8 | 2.8 | 2.8891 | 4.1 | 2.6 | < 5 | 3.39 | 3.5 | 2.1 | 1.8143 | < 5 | 3.3415 |  |  | 1
| 1 | 2 | 3.1 | 2.8725 | 2.5 | 0.1 | 18.29 | 2.8 | 2.2 | 2.4 | 2.5 | 3 | 2.9 | 2.9 | 2.7 | 2.7 | 2.7558 | 4 | 2.6 | < 5 | 3.4 | 3.2 | 2.3 | 1.8659 | < 5 | 3.1462 |  |  | 37
| 1 | 3 | 3.4 | 2.8554 | 2.5 | < 0.1 | 19.9706 | 3.1 | 2.9 | 3.5 | 2.7 | 2.9 | 2.9 | 2.8 | 2.7 | 3.1 | 2.81 | 4 | 2.5 | < 5 | 3.36 | 3.2 | 2.2 | 1.8219 | < 5 | 3.1942 |  |  | 16
| 1 | 4 | 3.1 | 2.8711 | 2.5 | < 0.1 | 17.7583 | 3 | 2.4 | 3.8 | 2.7 | 3 | 2.8 | 2.9 | 2.8 | 3.2 | 2.845 | 3.8 | 2.4 | < 5 | 3.38 | 3.3 | 2.3 | 1.8366 | < 5 | 3.1795 |  |  | 2.7974
| 1 | 5 | 2.9 | 2.7656 | 2.5 | 0.1 | 18.7933 | 2.9 | 2.6 | 2.3 | 2.6 | 3 | 2.7 | 2.9 | 2.7 | 3 | 2.8196 | 4 | 2.5 | < 5 | 3.3 | 3.4 | 2.2 | 1.8539 | < 5 | 3.1656 |  |  | 67
| 1 | 6 | 3 | 2.8661 | 2.5 | 0.2 | 18.11 | 2.8 | 2.6 | 3.1 | 2.9 | 3 | 2.8 | 2.9 | 2.7 | 3.1 | 2.7904 | 3.9 | 2.5 | < 5 | 3.13 | 3.3 | 2.2 | 1.8827 | < 5 | 3.1095
| Mean |  | 3.1333 | 2.8572 | 2.5 | 0.13333 | 18.7934 | 2.9 | 2.4833 | 3.9833 | 2.6333 | 2.9833 | 2.8167 | 2.85 | 2.7333 | 2.9833 | 2.8183 | 3.9667 | 2.5167 |  | 3.3267 | 3.3167 | 2.2167 | 1.8459 |  | 3.1894
| Median |  | 3.1 | 2.8686 | 2.5 | 0.1 | 18.5417 | 2.85 | 2.5 | 3.3 | 2.65 | 3 | 2.8 | 2.9 | 2.7 | 3.05 | 2.8148 | 4 | 2.5 |  | 3.37 | 3.3 | 2.2 | 1.8453 |  | 3.1725
| Std Dev. |  | 0.18619 | 0.048918 | 0 | 0.057735 | 0.924 | 0.12649 | 0.27142 | 2.4326 | 0.17512 | 0.040825 | 0.075277 | 0.083666 | 0.05164 | 0.19408 | 0.045776 | 0.10328 | 0.075277 |  | 0.1027 | 0.1169 | 0.075277 | 0.026379 |  | 0.080077
| Rel.Std.Dev. |  | 5.942% | 1.712% | 0% | 43.3% | 4.917% | 4.362% | 10.93% | 61.07% | 6.65% | 1.368% | 2.673% | 2.936% | 1.889% | 6.505% | 1.624% | 2.604% | 2.991% |  | 3.087% | 3.525% | 3.396% | 1.429% |  | 2.511%
| PDM3 |  | 12.01% | 2.137% | -10.63% | -95.23% | 571.8% | 3.666% | -11.23% | 42.39% | -5.866% | 6.645% | 0.6872% | 1.879% | -2.292% | 6.645% | 0.7459% | 41.8% | -10.04% |  | 18.92% | 18.56% | -20.76% | -34.02% |  | 14.01%
| Z-Score (Absolute) |  | 0.6 | 0.05 | 0.67 | 5.53 | 32.21 | 0.13 | 0.71 | 2.32 | 0.41 | 0.3 | 0.04 | 0.03 | 0.2 | 0.3 | 0.03 | 2.29 | 0.64 | 0.67 | 0.99 | 0.97 | 1.25 | 1.99 | 0.67 | 0.72
| Analytical results for Y in OREAS 238b (Certified Value 16 ppm) | [27 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.28 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 16.48 | 13.4884 | 17.7 | 19.7 | 10.52 | 18.6 | 14 | 18.5 | 14.5 | 18 | 17.2 | 16.7 | 16.3 | 17 | 15.5351 | 21.1 | 12.48 | 15.4 | 18.75 | 17 | 14.5 | 12.47 | 14.8321 | 23.4414 |  |  | 1
| 1 | 2 | 16.16 | 13.4433 | 18.2 | 20.7 | 11.02 | 17.8 | 14.4 | 18.3 | 14.3 | 18.5 | 16.9 | 17.2 | 16.6 | 14.1 | 15.4634 | 20.9 | 12.3 | 15 | 18.73 | 15.6 | 14.6 | 12.21 | 15.2426 | 23.234 |  |  | 38
| 1 | 3 | 16.77 | 13.4088 | 17.8 | 20.4 | 10.02 | 18.3 | 15.4 | 16.9 | 14.3 | 17.6 | 18.3 | 16.1 | 17.2 | 14.3 | 15.8006 | 21.3 | 12.75 | 13.9 | 22.73 | 15.3 | 14.7 | 12.55 | 15.5505 | 24.2894 |  |  | 16
| 1 | 4 | 16.59 | 13.5506 | 18.1 | 20.8 | 10.92 | 17.6 | 16.4 | 17 | 14.6 | 18.2 | 16.7 | 16.8 | 16.6 | 14.9 | 16.0221 | 20.9 | 12.6 | 15.3 | 19.02 | 15 | 14.6 | 12.4 | 14.3702 | 22.9307 |  |  | 16.0005
| 1 | 5 | 16.7 | 13.2663 | 18.2 | 20.3 | 11.035 | 18 | 16 | 16.4 | 14.9 | 17.2 | 16.6 | 16.3 | 17.2 | 14.4 | 15.887 | 20.7 | 12.42 | 14.4 | 20.31 | 16.3 | 14.9 | 12.44 | 14.165 | 22.2775 |  |  | 68
| 1 | 6 | 16.05 | 13.2305 | 17.9 | 20.6 | 11.705 | 18.4 | 15.4 | 17.4 | 14.3 | 17.7 | 16.7 | 16.8 | 16.8 | 13.8 | 16.1972 | 21.1 | 12.71 | 15.9 | 19.36 | 16 | 14.6 | 12.08 | 14.2163 | 23.4437
| Mean |  | 16.4583 | 13.398 | 17.9833 | 20.4167 | 10.87 | 18.1167 | 15.2667 | 17.4167 | 14.4833 | 17.8667 | 17.0667 | 16.65 | 16.7833 | 14.75 | 15.8176 | 21 | 12.5433 | 14.9833 | 19.8167 | 15.8667 | 14.65 | 12.3583 | 14.7294 | 23.2695
| Median |  | 16.535 | 13.4261 | 18 | 20.5 | 10.97 | 18.15 | 15.4 | 17.2 | 14.4 | 17.85 | 16.8 | 16.75 | 16.7 | 14.35 | 15.8438 | 21 | 12.54 | 15.15 | 19.19 | 15.8 | 14.6 | 12.42 | 14.6012 | 23.3377
| Std Dev. |  | 0.29294 | 0.12573 | 0.2137 | 0.39707 | 0.5647 | 0.38166 | 0.91797 | 0.82805 | 0.24014 | 0.46332 | 0.64083 | 0.3937 | 0.36009 | 1.1606 | 0.28152 | 0.20976 | 0.17443 | 0.72503 | 1.5423 | 0.72572 | 0.13784 | 0.17725 | 0.57602 | 0.66335
| Rel.Std.Dev. |  | 1.78% | 0.9384% | 1.188% | 1.945% | 5.195% | 2.107% | 6.013% | 4.754% | 1.658% | 2.593% | 3.755% | 2.365% | 2.146% | 7.868% | 1.78% | 0.9989% | 1.391% | 4.839% | 7.783% | 4.574% | 0.9409% | 1.434% | 3.911% | 2.851%
| PDM3 |  | 2.861% | -16.27% | 12.39% | 27.6% | -32.06% | 13.23% | -4.587% | 8.851% | -9.482% | 11.66% | 6.663% | 4.059% | 4.892% | -7.816% | -1.143% | 31.25% | -21.61% | -6.357% | 23.85% | -0.8367% | -8.441% | -22.76% | -7.944% | 45.43%
| Z-Score (Absolute) |  | 0.13 | 1.17 | 0.77 | 1.8 | 2.24 | 0.83 | 0.38 | 0.53 | 0.71 | 0.72 | 0.38 | 0.21 | 0.26 | 0.6 | 0.15 | 2.04 | 1.53 | 0.5 | 1.54 | 0.13 | 0.64 | 1.61 | 0.61 | 3
| Analytical results for Yb in OREAS 238b (Certified Value 1.82 ppm) | [27 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  | 1
| No | No | 1.01: 1.03; 1.03: 1.05; 1.05: 1.06; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.21; 1.10: 1.26; 1.11: 1.28; 1.12: 1.30; 1.13: 1.32; 1.14: 1.34 |  | ppm
|  |  | 1.01: 4A*MS; 1.03: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*MS; 1.09: 4A*MS; 1.10: 4A*MS; 1.11: 4A*OES/MS; 1.12: 4A*OES; 1.13: 4A*OES; 1.14: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 1.5602; 1.03: 1.95; 1.05: 2.2; 1.06: 2.9; 1.07: 2.08; 1.08: 1.49; 1.09: 1.5; 1.10: 2.1; 1.11: 1.5; 1.12: 1.31; 1.13: 1.6423; 1.14: 2.6067 |  | 1
| 1 | 2 | 1.01: 1.4953; 1.03: 1.9; 1.05: 2.3; 1.06: 1.5; 1.07: 2.09; 1.08: 1.459; 1.09: 1.5; 1.10: 2.07; 1.11: 1.5; 1.12: 1.29; 1.13: 1.6848; 1.14: 2.4186 |  | 39
| 1 | 3 | 1.01: 1.5125; 1.03: 1.85; 1.05: 2.2; 1.06: 2.5; 1.07: 2.09; 1.08: 1.516; 1.09: 1.5; 1.10: 2.07; 1.11: 1.5; 1.12: 1.39; 1.13: 1.7273; 1.14: 2.5672 |  | 16
| 1 | 4 | 1.01: 1.481; 1.03: 1.9; 1.05: 2.2; 1.06: 2; 1.07: 2.04; 1.08: 1.501; 1.09: 1.5; 1.10: 2.08; 1.11: 1.5; 1.12: 1.31; 1.13: 1.4724; 1.14: 2.46 |  | 1.8157
| 1 | 5 | 1.01: 1.5553; 1.03: 2; 1.05: 2.2; 1.06: 1.9; 1.07: 2.09; 1.08: 1.523; 1.09: 1.5; 1.10: 2.08; 1.11: 1.5; 1.12: 1.31; 1.13: 1.5574; 1.14: 2.4618 |  | 69
| 1 | 6 | 1.01: 1.5038; 1.03: 1.9; 1.05: 2.3; 1.06: 1.6; 1.07: 2.08; 1.08: 1.528; 1.09: 1.5; 1.10: 2; 1.11: 1.4; 1.12: 1.33; 1.13: 1.5574; 1.14: 2.4425
| Mean |  | 1.01: 1.518; 1.03: 1.9167; 1.05: 2.2333; 1.06: 2.0667; 1.07: 2.0783; 1.08: 1.5028; 1.09: 1.5; 1.10: 2.0667; 1.11: 1.4833; 1.12: 1.3233; 1.13: 1.6069; 1.14: 2.4928
| Median |  | 1.01: 1.5081; 1.03: 1.9; 1.05: 2.2; 1.06: 1.95; 1.07: 2.085; 1.08: 1.5085; 1.09: 1.5; 1.10: 2.075; 1.11: 1.5; 1.12: 1.31; 1.13: 1.5998; 1.14: 2.4609
| Std Dev. |  | 1.01: 0.032537; 1.03: 0.05164; 1.05: 0.05164; 1.06: 0.53914; 1.07: 0.019408; 1.08: 0.025701; 1.09: 0; 1.10: 0.034448; 1.11: 0.040825; 1.12: 0.035024; 1.13: 0.094666; 1.14: 0.075591
| Rel.Std.Dev. |  | 1.01: 2.143%; 1.03: 2.694%; 1.05: 2.312%; 1.06: 26.09%; 1.07: 0.9338%; 1.08: 1.71%; 1.09: 0%; 1.10: 1.667%; 1.11: 2.752%; 1.12: 2.647%; 1.13: 5.891%; 1.14: 3.032%
| PDM3 |  | 1.01: -16.4%; 1.03: 5.558%; 1.05: 23%; 1.06: 13.82%; 1.07: 14.46%; 1.08: -17.23%; 1.09: -17.39%; 1.10: 13.82%; 1.11: -18.31%; 1.12: -27.12%; 1.13: -11.5%; 1.14: 37.29%
| Z-Score (Absolute) |  | 1.01: 0.56; 1.03: 0.36; 1.05: 1.09; 1.06: 0.7; 1.07: 0.73; 1.08: 0.6; 1.09: 0.61; 1.10: 0.7; 1.11: 0.64; 1.12: 1.01; 1.13: 0.36; 1.14: 1.69
| Analytical results for Zn in OREAS 238b (Certified Value 100 ppm) | [27 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 0
1
| 1 | 1 | 101 | 95.7312 | 98 | 104 | 100.407 | 99 | 103 | 96 | 101 | 100 | 101 | 97 | 98 | 102.5 | 97.4388 | 103 | 99.63 | 99 | 111.1 | 98 | 98 | 97.3 | 100.078 | 105 | 89.2309 | 93.8679 | 1
| 1 | 2 | 100 | 96.4088 | 94 | 105 | 98.5171 | 99 | 100 | 98 | 102 | 100 | 104 | 94 | 97 | 99.6 | 98.1607 | 99 | 97.91 | 98 | 109.8 | 96 | 98 | 93.5 | 99.044 | 105 | 86.754 | 95.4196 | 40
| 1 | 3 | 104 | 97.4656 | 92 | 105 | 98.0883 | 100 | 104 | 97 | 111 | 100 | 104 | 95 | 96 | 100.7 | 98.7807 | 102 | 104.8 | 95 | 111.3 | 94 | 99 | 97.6 | 101.651 | 103 | 88.7529 | 96.5636 | 16
| 1 | 4 | 104 | 97.052 | 98 | 104 | 98.1502 | 99 | 106 | 98 | 108 | 102 | 101 | 95 | 99 | 100.8 | 97.6632 | 102 | 100.7 | 103 | 102.4 | 98 | 96 | 97.8 | 100.991 | 110 | 82.843 | 96.7899 | 99.9503
| 1 | 5 | 102 | 97.8136 | 96 | 106 | 97.57 | 100 | 100 | 96 | 107 | 100 | 100 | 96 | 101 | 100.7 | 97.3425 | 98 | 105 | 106 | 103.3 | 91 | 97 | 93.1 | 101.772 | 104 | 81.7277 | 96.4799 | 70
| 1 | 6 | 104 | 98.2192 | 98 | 105 | 99.6583 | 98 | 103 | 98 | 106 | 103 | 105 | 97 | 100 | 100.2 | 96.9465 | 102 | 99.14 | 95 | 101.1 | 97 | 98 | 99.8 | 99.924 | 100 | 81.2642 | 95.7635
| Mean |  | 102.5 | 97.1151 | 96 | 104.8333 | 98.7318 | 99.1667 | 102.6667 | 97.1667 | 105.8333 | 100.8333 | 102.5 | 95.6667 | 98.5 | 100.75 | 97.7221 | 101 | 101.1967 | 99.3333 | 106.5 | 95.6667 | 97.6667 | 96.5167 | 100.5767 | 104.5 | 85.0954 | 95.8141
| Median |  | 103 | 97.2588 | 97 | 105 | 98.3336 | 99 | 103 | 97.5 | 106.5 | 100 | 102.5 | 95.5 | 98.5 | 100.7 | 97.551 | 102 | 100.165 | 98.5 | 106.55 | 96.5 | 98 | 97.45 | 100.5345 | 104.5 | 84.7985 | 96.1217
| Std Dev. |  | 1.7607 | 0.92172 | 2.5298 | 0.75277 | 1.0782 | 0.75277 | 2.3381 | 0.98319 | 3.7639 | 1.3292 | 2.0736 | 1.2111 | 1.8708 | 0.96902 | 0.65459 | 2 | 3.0061 | 4.4121 | 4.7181 | 2.7325 | 1.0328 | 2.6453 | 1.075 | 3.2711 | 3.5867 | 1.0866
| Rel.Std.Dev. |  | 1.718% | 0.9491% | 2.635% | 0.7181% | 1.092% | 0.7591% | 2.277% | 1.012% | 3.556% | 1.318% | 2.023% | 1.266% | 1.899% | 0.9618% | 0.6698% | 1.98% | 2.971% | 4.442% | 4.43% | 2.856% | 1.057% | 2.741% | 1.069% | 3.13% | 4.215% | 1.134%
| PDM3 |  | 2.551% | -2.837% | -3.952% | 4.886% | -1.219% | -0.784% | 2.718% | -2.785% | 5.886% | 0.8835% | 2.551% | -4.286% | -1.451% | 0.8001% | -2.229% | 1.05% | 1.247% | -0.6172% | 6.553% | -4.286% | -2.285% | -3.435% | 0.6267% | 4.552% | -14.86% | -4.138%
| Z-Score (Absolute) |  | 0.9 | 0.59 | 0.9 | 1.55 | 0.14 | 0.02 | 0.95 | 0.58 | 1.82 | 0.44 | 0.9 | 0.99 | 0.21 | 0.42 | 0.42 | 0.48 | 0.54 | 0.02 | 2.01 | 0.99 | 0.44 | 0.76 | 0.37 | 1.45 | 3.92 | 0.95
| Analytical results for Zr in OREAS 238b (Certified Value 132 ppm) | [27 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 142.6 | 136.1318 | 152 | 161 | 111.49 | 130.8 | 119 | 135 | 133.3 | 138 | 137.5 | 134 | 133.5 | 131.9 | 127.4382 | 142 | 129 | 140.3 | 132.6 | 125.5 | 116.4 | 113.278 | 147 | 130.9817 | 220.7187 |  | 1
| 1 | 2 | 142.2 | 133.9954 | 146 | 175 | 113.12 | 129 | 125 | 139 | 132 | 138 | 135.5 | 133 | 136.5 | 124.9 | 129.1197 | 147 | 128 | 140.2 | 135.8 | 127.5 | 119 | 112.497 | 148 | 132.593 | 226.4067 |  | 41
| 1 | 3 | 143.7 | 136.6414 | 145 | 154 | 111.82 | 130.5 | 127 | 133 | 139.3 | 136 | 142.5 | 133 | 135.5 | 126 | 131.3984 | 167 | 127 | 142.5 | 135.1 | 127.4 | 117 | 114.389 | 145 | 136.3697 | 229.489 |  | 16
| 1 | 4 | 142.2 | 135.436 | 154 | 165 | 111.46 | 129.6 | 128 | 138 | 137 | 137.5 | 136.5 | 136 | 134 | 127.6 | 127.1951 | 143 | 127 | 145 | 139.1 | 127.5 | 118.5 | 112.629 | 150 | 108.2716 | 218.8973 |  | 132.1957
| 1 | 5 | 141.7 | 135.4458 | 142 | 164 | 114.98 | 128.8 | 130 | 126 | 139.1 | 136.5 | 136.5 | 130 | 137 | 127.1 | 131.6617 | 142 | 125 | 145.6 | 134.5 | 126 | 118.7 | 115.621 | 144 | 115.7745 | 223.8724 |  | 71
| 1 | 6 | 141.9 | 132.4176 | 143 | 166 | 112.72 | 128.8 | 126 | 131 | 140.1 | 137 | 141.5 | 137.5 | 138.5 | 123.3 | 128.3033 | 148 | 125 | 143.2 | 134.4 | 125.4 | 118.4 | 113.113 | 144 | 116.4291 | 227.0897
| Mean |  | 142.3833 | 135.0113 | 147 | 164.1667 | 112.5983 | 129.5833 | 125.8333 | 133.6667 | 136.8 | 137.1667 | 138.3333 | 133.9167 | 135.8333 | 126.8 | 129.1861 | 148.1667 | 126.8333 | 142.8 | 135.25 | 126.55 | 118 | 113.5878 | 146.3333 | 123.4033 | 224.4123
| Median |  | 142.2 | 135.4409 | 145.5 | 164.5 | 112.27 | 129.3 | 126.5 | 134 | 138.05 | 137.25 | 137 | 133.5 | 136 | 126.55 | 128.7115 | 145 | 127 | 142.85 | 134.8 | 126.7 | 118.45 | 113.1955 | 146 | 123.7054 | 225.1396
| Std Dev. |  | 0.71391 | 1.5519 | 4.899 | 6.8532 | 1.3483 | 0.88185 | 3.7639 | 4.8028 | 3.3988 | 0.8165 | 2.9269 | 2.6157 | 1.8886 | 2.9407 | 1.9404 | 9.5795 | 1.6021 | 2.2777 | 2.1659 | 1.0252 | 1.045 | 1.2005 | 2.4221 | 11.365 | 4.0316
| Rel.Std.Dev. |  | 0.5014% | 1.149% | 3.333% | 4.175% | 1.197% | 0.6805% | 2.991% | 3.593% | 2.485% | 0.5953% | 2.116% | 1.953% | 1.39% | 2.319% | 1.502% | 6.465% | 1.263% | 1.595% | 1.601% | 0.8101% | 0.8856% | 1.057% | 1.655% | 9.21% | 1.797%
| PDM3 |  | 7.707% | 2.13% | 11.2% | 24.18% | -14.82% | -1.976% | -4.813% | 1.113% | 3.483% | 3.76% | 4.643% | 1.302% | 2.752% | -4.082% | -2.277% | 12.08% | -4.056% | 8.022% | 2.31% | -4.271% | -10.74% | -14.08% | 10.69% | -6.651% | 69.76%
| Z-Score (Absolute) |  | 0.61 | 0 | 0.99 | 2.4 | 1.85 | 0.45 | 0.76 | 0.11 | 0.15 | 0.18 | 0.27 | 0.09 | 0.07 | 0.68 | 0.48 | 1.08 | 0.67 | 0.64 | 0.02 | 0.7 | 1.4 | 1.77 | 0.93 | 0.96 | 7.37
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C219: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C255: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C401: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C437: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C455: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C526: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C544: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1018: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1055: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1073: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1091: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1110: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1182: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1200: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1218: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1236: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 238b (Certified Value 0.24 ppm) | [25 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 0.23 | 0.26 | < 0.5 | 0.24 | < 0.5 | 0.26 | 0.2 | 0.24 | 0.25 | 0.24 | 0.24 | 0.224 | < 1 | 0.4 | 0.23 | 0.22 | < 0.5 | 0.25 | 0.26 | < 1 | 1.1941 |  |  |  | 1
| 1 | 2 | 0.23 | 0.248 | < 0.5 | 0.25 | < 0.5 | 0.26 | 0.2 | 0.23 | 0.24 | 0.23 | 0.24 | 0.206 | < 1 | 0.3 | 0.21 | 0.24 | < 0.5 | 0.23 | 0.26 | < 1 | 1.2937 |  |  |  | 25
| 1 | 3 | 0.23 | 0.24 | < 0.5 | 0.24 | < 0.5 | 0.26 | 0.2 | 0.24 | 0.25 | 0.22 | 0.23 | 0.224 | < 1 | 0.3 | 0.25 | 0.2 | < 0.5 | 0.24 | 0.25 | < 1 | 1.3601 |  |  |  | 16
| 1 | 4 | 0.23 | 0.247 | < 0.5 | 0.22 | < 0.5 | 0.25 | 0.2 | 0.24 | 0.25 | 0.24 | 0.25 | 0.228 | < 1 | 0.3 | 0.24 | 0.12 | < 0.5 | 0.23 | 0.26 | < 1 | 1.3321 |  |  |  | 0.23976
| 1 | 5 | 0.23 | 0.254 | < 0.5 | 0.24 | < 0.5 | 0.25 | 0.2 | 0.24 | 0.24 | 0.24 | 0.23 | 0.209 | < 1 | 0.4 | 0.24 | 0.13 | < 0.5 | 0.23 | 0.26 | < 1 | 1.3015 |  |  |  | 73
| 1 | 6 | 0.22 | 0.242 | < 0.5 | 0.22 | < 0.5 | 0.27 | 0.3 | 0.25 | 0.24 | 0.24 | 0.24 | 0.211 | < 1 | 0.3 | 0.24 | 0.15 | < 0.5 | 0.24 | 0.24 | < 1 | 1.162
| Mean |  | 0.22833 | 0.2485 |  | 0.235 |  | 0.25833 | 0.21667 | 0.24 | 0.245 | 0.235 | 0.23833 | 0.217 |  | 0.33333 | 0.235 | 0.17667 |  | 0.23667 | 0.255 |  | 1.2739
| Median |  | 0.23 | 0.2475 |  | 0.24 |  | 0.26 | 0.2 | 0.24 | 0.245 | 0.24 | 0.24 | 0.2175 |  | 0.3 | 0.24 | 0.175 |  | 0.235 | 0.26 |  | 1.2976
| Std Dev. |  | 0.0040825 | 0.0074766 |  | 0.012247 |  | 0.0075277 | 0.040825 | 0.0063246 | 0.0054772 | 0.0083666 | 0.0075277 | 0.0093808 |  | 0.05164 | 0.013784 | 0.050067 |  | 0.008165 | 0.0083666 |  | 0.078564
| Rel.Std.Dev. |  | 1.788% | 3.009% |  | 5.212% |  | 2.914% | 18.84% | 2.635% | 2.236% | 3.56% | 3.158% | 4.323% |  | 15.49% | 5.866% | 28.34% |  | 3.45% | 3.281% |  | 6.167%
| PDM3 |  | -4.767% | 3.644% |  | -1.987% |  | 7.745% | -9.633% | 0.09848% | 2.184% | -1.987% | -0.5967% | -9.494% |  | 39.03% | -1.987% | -26.32% |  | -1.292% | 6.355% |  | 431.3%
| Z-Score (Absolute) |  | 1.12 | 0.24 | 0.34 | 0.67 | 0.34 | 0.9 | NA | 0.34 | 0 | 0.67 | 0.45 | 1.89 | 17.19 | NA | 0.67 | 4.61 | 0.34 | 0.56 | 0.67 | 17.19 | 69.38
Results from laboratories 1.11 and 1.19 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 238b (Certified Value 2.62 wt.%) | [25 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 2.433 | 3.0379 | 2.71 | 2.8633 | 2.797 | 2.56 | 3.03 | 2.51 | 2.49 | 2.62 | 2.6 | 2.59 | 2.53 | 2.5833 | 2.466 | 2.49 | 2.6 | 2.5 | 2.3989 | 4.9799 |  |  |  |  | 1
| 1 | 2 | 2.428 | 3.0755 | 2.76 | 2.85 | 2.679 | 2.55 | 3.12 | 2.51 | 2.5 | 2.68 | 2.57 | 2.63 | 2.57 | 2.5393 | 2.419 | 2.39 | 2.6 | 2.41 | 2.3223 | 4.9813 |  |  |  |  | 0
| 1 | 3 | 2.44 | 3.0813 | 2.75 | 2.8367 | 2.785 | 2.59 | 3.1 | 2.48 | 2.45 | 2.73 | 2.59 | 2.63 | 2.61 | 2.5714 | 2.472 | 2.49 | 2.55 | 2.44 | 2.2806 | 4.9773 |  |  |  |  | 16
| 1 | 4 | 2.426 | 3.0876 | 2.42 | 2.8567 | 2.721 | 2.52 | 3.1 | 2.52 | 2.49 | 2.74 | 2.6 | 2.65 | 2.67 | 2.6037 | 2.366 | 2.44 | 2.56 | 2.37 | 2.3434 | 4.9867 |  |  |  |  | 2.6181
| 1 | 5 | 2.418 | 3.0603 | 2.64 | 2.83 | 2.677 | 2.6 | 3.07 | 2.49 | 2.53 | 2.78 | 2.61 | 2.62 | 2.53 | 2.603 | 2.429 | 2.36 | 2.6 | 2.32 | 2.3524 | 4.9847 |  |  |  |  | 74
| 1 | 6 | 2.411 | 3.0267 | 2.76 | 2.87 | 2.715 | 2.62 | 3.14 | 2.54 | 2.51 | 2.75 | 2.56 | 2.65 | 2.48 | 2.6274 | 2.435 | 2.48 | 2.59 | 2.37 | 2.3316 | 4.9842
| Mean |  | 2.426 | 3.0615 | 2.6733 | 2.8511 | 2.729 | 2.5733 | 3.0933 | 2.5083 | 2.495 | 2.7167 | 2.5883 | 2.6283 | 2.565 | 2.588 | 2.4312 | 2.4417 | 2.5833 | 2.4017 | 2.3382 | 4.9824
| Median |  | 2.427 | 3.0679 | 2.73 | 2.8533 | 2.718 | 2.575 | 3.1 | 2.51 | 2.495 | 2.735 | 2.595 | 2.63 | 2.55 | 2.5932 | 2.432 | 2.46 | 2.595 | 2.39 | 2.3375 | 4.9828
| Std Dev. |  | 0.010373 | 0.024652 | 0.13231 | 0.015444 | 0.051428 | 0.036697 | 0.038816 | 0.02137 | 0.026646 | 0.057504 | 0.019408 | 0.022286 | 0.067454 | 0.030648 | 0.038196 | 0.055648 | 0.022509 | 0.063061 | 0.0388 | 0.0034778
| Rel.Std.Dev. |  | 0.4276% | 0.8052% | 4.949% | 0.5417% | 1.884% | 1.426% | 1.255% | 0.852% | 1.068% | 2.117% | 0.7498% | 0.8479% | 2.63% | 1.184% | 1.571% | 2.279% | 0.8713% | 2.626% | 1.659% | 0.0698%
| PDM3 |  | -7.337% | 16.94% | 2.11% | 8.9% | 4.236% | -1.71% | 18.15% | -4.193% | -4.702% | 3.765% | -1.137% | 0.391% | -2.028% | -1.149% | -7.14% | -6.739% | -1.328% | -8.267% | -10.69% | 90.3%
| Z-Score (Absolute) |  | 0.79 | 2.34 | 0.43 | 1.3 | 0.7 | 0.06 | 2.5 | 0.38 | 0.45 | 0.64 | 0.01 | 0.21 | 0.1 | 0.01 | 0.76 | 0.71 | 0.01 | 0.9 | 1.22 | 11.78
| Analytical results for As in OREAS 238b (Certified Value 674 ppm) | [25 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g |  | 0
| 1 | 1 | 674.57 | 697.89 | 635 | 659 | 671 | 701.6 | 610 | 695.5 | 679 | 692 | 706 | 695 | 698.4 | 681.4665 | 662 | 600 | 772.64 | 684.7 | 647.7 | 661 | 646 | 671.0616 | 660.3564 |  | 1
| 1 | 2 | 683.89 | 704.47 | 627 | 667.1292 | 669 | 695.3 | 642 | 703.6 | 679 | 689 | 685 | 695 | 690.5 | 697.5278 | 644 | 564 | 712.44 | 682.8 | 624.2 | 643.4 | 664 | 690.108 | 667.4631 |  | 26
| 1 | 3 | 687.47 | 698.75 | 633 | 662.9186 | 676 | 689.9 | 628 | 709.5 | 704 | 698 | 688 | 703 | 693 | 686.3359 | 663 | 608 | 749.23 | 686.4 | 637.2 | 623.7 | 658 | 689.4534 | 652.1058 |  | 16
| 1 | 4 | 683.04 | 706.52 | 595 | 666.2713 | 669 | 687.2 | 626 | 725.5 | 688 | 694 | 685 | 706 | 708.5 | 671.6813 | 635 | 608 | 781.66 | 686.2 | 629.9 | 641.5 | 694 | 393.2365 | 657.9708 |  | 673.6185
| 1 | 5 | 687.9 | 695.9 | 640 | 672.14 | 678 | 691.9 | 643 | 702.7 | 678 | 702 | 701 | 698 | 676.6 | 671.7326 | 651 | 599 | 745.33 | 683.4 | 620.9 | 642.5 | 681 | 623.489 | 655.9753 |  | 75
| 1 | 6 | 672.9 | 701.73 | 641 | 665.5708 | 675 | 687.1 | 649 | 729 | 689 | 698 | 695 | 701 | 669.6 | 691.5052 | 658 | 626 | 751.84 | 688.1 | 624.8 | 639.7 | 689 | 654.3359 | 655.9102
| Mean |  | 681.6283 | 700.8767 | 628.5 | 665.505 | 673 | 692.1667 | 633 | 710.9667 | 686.1667 | 695.5 | 693.3333 | 699.6667 | 689.4333 | 683.3749 | 652.1667 | 600.8333 | 752.19 | 685.2667 | 630.7833 | 641.9667 | 672 | 620.2807 | 658.2969
| Median |  | 683.465 | 700.24 | 634 | 665.9211 | 673 | 690.9 | 635 | 706.55 | 683.5 | 696 | 691.5 | 699.5 | 691.75 | 683.9012 | 654.5 | 604 | 750.535 | 685.45 | 627.35 | 642 | 672.5 | 662.6988 | 656.973
| Std Dev. |  | 6.4276 | 4.0914 | 17.1785 | 4.3877 | 3.8471 | 5.5561 | 14.4222 | 13.4193 | 9.9883 | 4.7223 | 8.8242 | 4.4572 | 14.2538 | 10.4986 | 11.0529 | 20.4784 | 24.1772 | 2.0037 | 10.0494 | 11.8628 | 18.9209 | 113.9812 | 5.2494
| Rel.Std.Dev. |  | 0.943% | 0.5838% | 2.733% | 0.6593% | 0.5716% | 0.8027% | 2.278% | 1.887% | 1.456% | 0.679% | 1.273% | 0.637% | 2.067% | 1.536% | 1.695% | 3.408% | 3.214% | 0.2924% | 1.593% | 1.848% | 2.816% | 18.38% | 0.7974%
| PDM3 |  | 1.189% | 4.047% | -6.698% | -1.204% | -0.09181% | 2.754% | -6.03% | 5.544% | 1.863% | 3.248% | 2.927% | 3.867% | 2.348% | 1.448% | -3.185% | -10.81% | 11.66% | 1.729% | -6.359% | -4.699% | -0.2403% | -7.918% | -2.275%
| Z-Score (Absolute) |  | 0 | 0.72 | 1.99 | 0.6 | 0.32 | 0.39 | 1.82 | 1.1 | 0.17 | 0.52 | 0.44 | 0.67 | 0.29 | 0.07 | 1.1 | 3.02 | 2.64 | 0.14 | 1.9 | 1.48 | 0.36 | 2.29 | 0.87
| Analytical results for B in OREAS 238b (Indicative Value 15.5 ppm) | [25 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.06; 1.07: 1.07; 1.08: 1.08; 1.09: 1.12; 1.10: 1.13; 1.11: 1.14; 1.12: 1.15; 1.13: 1.16; 1.14: 1.27; 1.15: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES; 1.08: AR*OES/MS; 1.09: AR*MS; 1.10: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*OES/MS; 1.15: AR*OES |  |  |  | 1
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 01g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g; 1.14: 0.5g; 1.15: 0.5g |  |  |  | 1
| 1 | 1 | 1.01: < 5; 1.06: 5; 1.07: 28.72; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21 |  |  |  | 1
| 1 | 2 | 1.01: < 5; 1.06: 5; 1.07: 26.86; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 20 |  |  |  | 4
| 1 | 3 | 1.01: < 5; 1.06: 5; 1.07: 26.11; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 20 |  |  |  | 16
| 1 | 4 | 1.01: < 5; 1.06: 5; 1.07: 26.345; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 20 |  |  |  | 15.509
| 1 | 5 | 1.01: < 5; 1.06: 5; 1.07: 24.54; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21 |  |  |  | 14
| 1 | 6 | 1.01: < 5; 1.06: 5; 1.07: 26.64; 1.08: < 10; 1.09: < 10; 1.10: 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10; 1.15: 21
| Mean |  | 1.06: 5; 1.07: 26.5358; 1.10: 10; 1.15: 20.5
| Median |  | 1.06: 5; 1.07: 26.4925; 1.10: 10; 1.15: 20.5
| Std Dev. |  | 1.06: 0; 1.07: 1.3482; 1.10: 0; 1.15: 0.54772
| Rel.Std.Dev. |  | 1.06: 0%; 1.07: 5.081%; 1.10: 0%; 1.15: 2.672%
| PDM3 |  | 1.06: -67.76%; 1.07: 71.1%; 1.10: -35.52%; 1.15: 32.18%
Z-Score (Absolute) | bar: 1.01 to 1.15
| Analytical results for Ba in OREAS 238b (Certified Value 103 ppm) | [25 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 96.47 | 130.2806 | 88 | 109.3133 | 107 | 105 | 131 | 102 | 100 | 90 | 110 | 100 | 99.5 | 104.6601 | 96 | 107.9 | 81 | 111.8 | 112 | 107 | 101.48 | 594.0052 |  |  | 1
| 1 | 2 | 96.37 | 129.6172 | 73.2 | 108.0633 | 105 | 99 | 134 | 100 | 100 | 90 | 100 | 100 | 101.3 | 99.5459 | 95 | 120.8 | 78 | 103.71 | 113 | 103 | 97.41 | 599.79 |  |  | 27
| 1 | 3 | 97.02 | 132.3765 | 98.5 | 108.3767 | 105 | 100 | 127 | 99 | 90 | 90 | 100 | 100 | 100.7 | 101.1355 | 99 | 113.6 | 81 | 103.09 | 110.4 | 105 | 95.44 | 610.7063 |  |  | 16
| 1 | 4 | 96.11 | 131.9085 | 70.3 | 107.1067 | 107 | 99 | 128 | 103 | 100 | 100 | 100 | 100 | 103.5 | 100.881 | 94 | 108.6 | 79 | 105.21 | 111.7 | 102 | 99.08 | 594.4995 |  |  | 102.5083
| 1 | 5 | 96.26 | 130.1071 | 116 | 106.895 | 104 | 102 | 123 | 100 | 100 | 100 | 110 | 100 | 98.9 | 102.0302 | 95 | 110.1 | 77 | 106.19 | 111.9 | 100 | 98.61 | 595.5603 |  |  | 76
| 1 | 6 | 95.38 | 132.1963 | 114 | 108.205 | 104 | 103 | 133 | 107 | 100 | 100 | 100 | 100 | 96.7 | 102.3221 | 96 | 119.8 | 81 | 113.89 | 110.8 | 102 | 97.89 | 595.9328
| Mean |  | 96.2683 | 131.081 | 93.3333 | 107.9933 | 105.3333 | 101.3333 | 129.3333 | 101.8333 | 98.3333 | 95 | 103.3333 | 100 | 100.1 | 101.7625 | 95.8333 | 113.4667 | 79.5 | 107.315 | 111.6333 | 103.1667 | 98.3183 | 598.4157
| Median |  | 96.315 | 131.0946 | 93.25 | 108.1342 | 105 | 101 | 129.5 | 101 | 100 | 95 | 100 | 100 | 100.1 | 101.5829 | 95.5 | 111.85 | 80 | 105.7 | 111.8 | 102.5 | 98.25 | 595.7465
| Std Dev. |  | 0.53484 | 1.2115 | 19.6566 | 0.88674 | 1.3663 | 2.4221 | 4.1312 | 2.9269 | 4.0825 | 5.4772 | 5.164 | 0 | 2.3117 | 1.7242 | 1.7224 | 5.6553 | 1.7607 | 4.4697 | 0.92664 | 2.4833 | 1.9977 | 6.357
| Rel.Std.Dev. |  | 0.5556% | 0.9242% | 21.06% | 0.8211% | 1.297% | 2.39% | 3.194% | 2.874% | 4.152% | 5.766% | 4.997% | 0% | 2.309% | 1.694% | 1.797% | 4.984% | 2.215% | 4.165% | 0.8301% | 2.407% | 2.032% | 1.062%
| PDM3 |  | -6.087% | 27.87% | -8.95% | 5.351% | 2.756% | -1.146% | 26.17% | -0.6584% | -4.073% | -7.325% | 0.8049% | -2.447% | -2.349% | -0.7276% | -6.512% | 10.69% | -22.45% | 4.689% | 8.902% | 0.6423% | -4.087% | 483.8%
| Z-Score (Absolute) |  | 0.68 | 3.57 | 1.03 | 0.76 | 0.43 | 0.06 | 3.36 | 0 | 0.42 | 0.83 | 0.19 | 0.22 | 0.21 | 0 | 0.73 | 1.42 | 2.72 | 0.67 | 1.2 | 0.17 | 0.42 | 60.63
| Analytical results for Be in OREAS 238b (Certified Value 1.17 ppm) | [25 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.11 | 1.5658 | 1 | < 5 | 1.1 | 1.16 | 1.33 | 1 | 1.26 | 1.14 | 1.21 | 1.21 | 1.2114 | 1.2 | 1.217 | 1.2 | 1.15 | < 2 | 1.03 | 0.948 | 1.1034 | 1.1737 |  |  | 1
| 1 | 2 | 1.14 | 1.5287 | 1.1 | < 5 | 1.1 | 1.16 | 1.39 | < 1 | 1.24 | 1.11 | 1.24 | 1.21 | 1.3088 | 1.2 | 1.282 | 1.1 | 1.1 | < 2 | 1 | 0.92 | 1.1685 | 1.2074 |  |  | 28
| 1 | 3 | 1.07 | 1.358 | 1 | < 5 | 1 | 1.16 | 1.27 | < 1 | 1.26 | 1.18 | 1.23 | 1.11 | 1.1944 | 1.3 | 1.232 | 1.2 | 1.16 | < 2 | 1.03 | 0.901 | 1.1469 | 1.1719 |  |  | 16
| 1 | 4 | 1.12 | 1.5072 | 1.1 | < 5 | 1.1 | 1.13 | 1.28 | 1 | 1.36 | 1.21 | 1.21 | 1.22 | 1.2153 | 1.2 | 1.178 | 1.1 | 1.19 | < 2 | 0.96 | 0.925 | 1.0817 | 1.099 |  |  | 1.169
| 1 | 5 | 1.15 | 1.3909 | 1.1 | < 5 | 1.1 | 1.15 | 1.27 | 1 | 1.3 | 1.12 | 1.22 | 1.21 | 1.298 | 1.2 | 1.152 | 1.1 | 1.18 | < 2 | 0.96 | 0.93 | 1.1249 | 1.1597 |  |  | 77
| 1 | 6 | 1.12 | 1.5529 | 1.1 | < 5 | 1.1 | 1.17 | 1.41 | 1 | 1.3 | 1.16 | 1.15 | 1.17 | 1.2447 | 1.2 | 1.232 | 1.1 | 1.2 | < 2 | 0.97 | 0.927 | 1.1248 | 1.1979
| Mean |  | 1.1183 | 1.4839 | 1.0667 |  | 1.0833 | 1.155 | 1.325 | 1 | 1.2867 | 1.1533 | 1.21 | 1.1883 | 1.2454 | 1.2167 | 1.2155 | 1.1333 | 1.1633 |  | 0.99167 | 0.92517 | 1.125 | 1.1683
| Median |  | 1.12 | 1.5179 | 1.1 |  | 1.1 | 1.16 | 1.305 | 1 | 1.28 | 1.15 | 1.215 | 1.21 | 1.23 | 1.2 | 1.2245 | 1.1 | 1.17 |  | 0.985 | 0.926 | 1.1248 | 1.1728
| Std Dev. |  | 0.027869 | 0.087769 | 0.05164 |  | 0.040825 | 0.013784 | 0.06253 | 0 | 0.043205 | 0.037771 | 0.031623 | 0.042151 | 0.047863 | 0.040825 | 0.045623 | 0.05164 | 0.036148 |  | 0.033116 | 0.015224 | 0.030703 | 0.038262
| Rel.Std.Dev. |  | 2.492% | 5.915% | 4.841% |  | 3.768% | 1.193% | 4.719% | 0% | 3.358% | 3.275% | 2.613% | 3.547% | 3.843% | 3.355% | 3.753% | 4.556% | 3.107% |  | 3.339% | 1.646% | 2.729% | 3.275%
| PDM3 |  | -4.337% | 26.93% | -8.757% |  | -7.331% | -1.201% | 13.34% | -14.46% | 10.06% | -1.343% | 3.504% | 1.651% | 6.535% | 4.075% | 3.975% | -3.054% | -0.4877% |  | -15.17% | -20.86% | -3.764% | -0.06409%
| Z-Score (Absolute) |  | 0.57 | 4.05 | 1.22 | 16.9 | 1.01 | 0.11 | 2.04 | NA | 1.56 | 0.13 | 0.59 | 0.32 | 1.04 | 0.67 | 0.66 | 0.38 | 0 | 2.07 | 2.17 | 3.01 | 0.48 | 0.06
Results from laboratory 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 238b (Certified Value 0.38 ppm) | [25 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.37 | 0.39524 | 0.4 | < 2 | 0.45 | 0.4 | 0.37 | 0.4 | 0.36 | 0.36 | 0.37 | 0.39 | 0.35 | 0.38941 | 2 | 0.519 | 0.34 | 42.17 | < 5 | 0.36 | 0.5544 | 1.3205 |  |  | 1
| 1 | 2 | 0.36 | 0.39406 | 0.37 | < 2 | 0.45 | 0.36 | 0.38 | 0.4 | 0.35 | 0.36 | 0.37 | 0.43 | 0.34 | 0.40224 | 3 | 0.688 | 0.36 | 39.23 | < 5 | 0.34 | 0.488 | 1.9432 |  |  | 29
| 1 | 3 | 0.35 | 0.41665 | 0.4 | < 2 | 0.46 | 0.4 | 0.41 | 0.4 | 0.35 | 0.35 | 0.35 | 0.39 | 0.38 | 0.38133 | 3 | 0.494 | 0.37 | 41.51 | < 5 | 0.36 | 0.4477 | 1.3288 |  |  | 16
| 1 | 4 | 0.35 | 0.42032 | 0.33 | < 2 | 0.45 | 0.35 | 0.39 | 0.4 | 0.35 | 0.36 | 0.34 | 0.4 | 0.37 | 0.36351 | 4 | 0.441 | 0.35 | 31.27 | < 5 | 0.34 | 0.5293 | 1.2483 |  |  | 0.37697
| 1 | 5 | 0.37 | 0.3852 | 0.4 | < 2 | 0.45 | 0.36 | 0.38 | 0.3 | 0.36 | 0.39 | 0.35 | 0.41 | 0.32 | 0.36549 | 4 | 0.479 | 0.35 | 29.23 | < 5 | 0.33 | 0.5441 | 1.6196 |  |  | 78
| 1 | 6 | 0.36 | 0.41323 | 0.4 | < 2 | 0.46 | 0.42 | 0.38 | 0.4 | 0.37 | 0.35 | 0.35 | 0.38 | 0.37 | 0.38847 | 4 | 0.495 | 0.36 | 23.03 | < 5 | 0.34 | 0.4393 | 1.4125
| Mean |  | 0.36 | 0.40412 | 0.38333 |  | 0.45333 | 0.38167 | 0.385 | 0.38333 | 0.35667 | 0.36167 | 0.355 | 0.4 | 0.355 | 0.38174 | 3.3333 | 0.51933 | 0.355 | 34.4067 |  | 0.345 | 0.50047 | 1.4788
| Median |  | 0.36 | 0.40423 | 0.4 |  | 0.45 | 0.38 | 0.38 | 0.4 | 0.355 | 0.36 | 0.35 | 0.395 | 0.36 | 0.3849 | 3.5 | 0.4945 | 0.355 | 35.25 |  | 0.34 | 0.50865 | 1.3706
| Std Dev. |  | 0.0089443 | 0.014427 | 0.028752 |  | 0.005164 | 0.028577 | 0.013784 | 0.040825 | 0.008165 | 0.01472 | 0.012247 | 0.017889 | 0.022583 | 0.014974 | 0.8165 | 0.086539 | 0.010488 | 7.7467 |  | 0.012247 | 0.049649 | 0.26099
| Rel.Std.Dev. |  | 2.485% | 3.57% | 7.5% |  | 1.139% | 7.488% | 3.58% | 10.65% | 2.289% | 4.07% | 3.45% | 4.472% | 6.361% | 3.923% | 24.49% | 16.66% | 2.954% | 22.52% |  | 3.55% | 9.921% | 17.65%
| PDM3 |  | -4.501% | 7.203% | 1.689% |  | 20.26% | 1.247% | 2.131% | 1.689% | -5.385% | -4.059% | -5.827% | 6.11% | -5.827% | 1.267% | 784.3% | 37.77% | -5.827% | 9027% |  | -8.48% | 32.76% | 292.3%
| Z-Score (Absolute) |  | 0.56 | 0.46 | 0.02 | 14.24 | 1.6 | 0.06 | 0.02 | NA | 0.64 | 0.52 | 0.67 | 0.37 | 0.67 | 0.06 | NA | 3.12 | 0.67 | 786.57 | 48.92 | 0.91 | 2.69 | 25.31
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 238b (Certified Value 0.309 wt.%) | [25 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.297 | 0.31276 | 0.36 | 0.35767 | 0.32 | 0.31 | 0.31 | 0.33 | 0.29 | 0.31 | 0.32 | 0.31 | 0.3 | 0.33611 | 0.3 | 0.28 | 0.31399 | 0.3 | 0.31 | 0.2741 |  |  |  |  | 1
| 1 | 2 | 0.297 | 0.31455 | 0.36 | 0.35867 | 0.32 | 0.3 | 0.32 | 0.33 | 0.29 | 0.31 | 0.32 | 0.31 | 0.31 | 0.32571 | 0.293 | 0.27 | 0.3169 | 0.31 | 0.3 | 0.2679 |  |  |  |  | 0
| 1 | 3 | 0.3 | 0.31817 | 0.37 | 0.34967 | 0.32 | 0.31 | 0.32 | 0.32 | 0.29 | 0.32 | 0.32 | 0.31 | 0.31 | 0.3353 | 0.3 | 0.28 | 0.31545 | 0.3 | 0.3 | 0.2589 |  |  |  |  | 16
| 1 | 4 | 0.301 | 0.31747 | 0.34 | 0.359 | 0.32 | 0.3 | 0.32 | 0.33 | 0.29 | 0.32 | 0.32 | 0.31 | 0.32 | 0.33155 | 0.286 | 0.28 | 0.31266 | 0.28 | 0.3 | 0.2672 |  |  |  |  | 0.30901
| 1 | 5 | 0.301 | 0.31495 | 0.38 | 0.34 | 0.32 | 0.31 | 0.32 | 0.33 | 0.29 | 0.32 | 0.32 | 0.31 | 0.3 | 0.33224 | 0.293 | 0.27 | 0.31819 | 0.3 | 0.29 | 0.2669 |  |  |  |  | 79
| 1 | 6 | 0.301 | 0.31257 | 0.37 | 0.36233 | 0.32 | 0.3 | 0.32 | 0.33 | 0.3 | 0.32 | 0.32 | 0.31 | 0.3 | 0.32908 | 0.3 | 0.28 | 0.32642 | 0.29 | 0.3 | 0.2646
| Mean |  | 0.2995 | 0.31508 | 0.36333 | 0.35456 | 0.32 | 0.305 | 0.31833 | 0.32833 | 0.29167 | 0.31667 | 0.32 | 0.31 | 0.30667 | 0.33166 | 0.29533 | 0.27667 | 0.31727 | 0.29667 | 0.3 | 0.2666
| Median |  | 0.3005 | 0.31475 | 0.365 | 0.35817 | 0.32 | 0.305 | 0.32 | 0.33 | 0.29 | 0.32 | 0.32 | 0.31 | 0.305 | 0.33189 | 0.2965 | 0.28 | 0.31617 | 0.3 | 0.3 | 0.26705
| Std Dev. |  | 0.0019748 | 0.0023321 | 0.013663 | 0.0082776 | 0 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0040825 | 0.005164 | 0 | 0 | 0.008165 | 0.003882 | 0.0057155 | 0.005164 | 0.0048999 | 0.010328 | 0.0063246 | 0.0049332
| Rel.Std.Dev. |  | 0.6594% | 0.7402% | 3.76% | 2.335% | 0% | 1.796% | 1.282% | 1.243% | 1.4% | 1.631% | 0% | 0% | 2.662% | 1.17% | 1.935% | 1.866% | 1.544% | 3.481% | 2.108% | 1.85%
| PDM3 |  | -3.077% | 1.965% | 17.58% | 14.74% | 3.557% | -1.297% | 3.018% | 6.254% | -5.612% | 2.478% | 3.557% | 0.321% | -0.7578% | 7.332% | -4.425% | -10.47% | 2.673% | -3.994% | -2.915% | -13.72%
| Z-Score (Absolute) |  | 0.69 | 0.13 | 2.68 | 2.22 | 0.39 | 0.4 | 0.31 | 0.83 | 1.1 | 0.22 | 0.39 | 0.13 | 0.31 | 1.01 | 0.9 | 1.89 | 0.25 | 0.84 | 0.66 | 2.42
| Analytical results for Cd in OREAS 238b (Certified Value 0.08 ppm) | [25 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.062 | 1.2 | 0.09 | 0.6609 | 0.11 | 0.08 | 0.1 | < 0.1 | 0.06 | 0.06 | 0.07 | 0.07 | 0.08 | 0.08386 | < 1 | < 0.1 | 0.11 | 24.12 | 12.7 | 0.07 | 1.01 | 0.3688 |  |  | 1
| 1 | 2 | 0.068 | 1.21 | 0.07 | 0.65203 | 0.1 | 0.08 | 0.1 | < 0.1 | 0.09 | 0.07 | 0.05 | 0.07 | 0.07 | 0.082552 | < 1 | 0.108 | 0.08 | 22.34 | 12.5 | 0.07 | 0.76 | 0.3434 |  |  | 30
| 1 | 3 | 0.07 | 1.25 | 0.06 | 0.67 | 0.12 | 0.09 | 0.11 | < 0.1 | 0.06 | 0.06 | 0.06 | 0.07 | 0.08 | 0.081273 | < 1 | < 0.1 | 0.08 | 22.64 | 12.2 | 0.07 | 0.81 | 0.3383 |  |  | 16
| 1 | 4 | 0.062 | 1.28 | 0.06 | 0.64653 | 0.11 | 0.08 | 0.1 | 0.1 | 0.07 | 0.07 | 0.06 | 0.07 | 0.08 | 0.081471 | < 1 | < 0.1 | 0.08 | 23.36 | 12.8 | 0.1 | 0.88 | 0.3357 |  |  | 0.080003
| 1 | 5 | 0.067 | 1.3 | 0.1 | 0.65 | 0.12 | 0.09 | 0.1 | < 0.1 | 0.1 | 0.07 | 0.06 | 0.07 | 0.07 | 0.08402 | < 1 | < 0.1 | 0.07 | 22.77 | 12.5 | 0.06 | 0.85 | 0.3382 |  |  | 80
| 1 | 6 | 0.066 | 1.22 | 0.05 | 0.6718 | 0.11 | 0.09 | 0.1 | 0.2 | 0.06 | 0.07 | 0.06 | 0.07 | 0.08 | 0.082054 | < 1 | 0.114 | 0.07 | 23.43 | 12.5 | 0.07 | 0.88 | 0.3382
| Mean |  | 0.065833 | 1.2433 | 0.071667 | 0.65854 | 0.11167 | 0.085 | 0.10167 | 0.15 | 0.073333 | 0.066667 | 0.06 | 0.07 | 0.076667 | 0.082538 |  | 0.111 | 0.081667 | 23.11 | 12.5333 | 0.073333 | 0.865 | 0.34377
| Median |  | 0.0665 | 1.235 | 0.065 | 0.65647 | 0.11 | 0.085 | 0.1 | 0.15 | 0.065 | 0.07 | 0.06 | 0.07 | 0.08 | 0.082303 |  | 0.111 | 0.08 | 23.065 | 12.5 | 0.07 | 0.865 | 0.33825
| Std Dev. |  | 0.0032506 | 0.040332 | 0.019408 | 0.010697 | 0.0075277 | 0.0054772 | 0.0040825 | 0.070711 | 0.017512 | 0.005164 | 0.0063246 | 0 | 0.005164 | 0.0011762 |  | 0.0042426 | 0.01472 | 0.65029 | 0.20656 | 0.013663 | 0.084558 | 0.01252
| Rel.Std.Dev. |  | 4.938% | 3.244% | 27.08% | 1.624% | 6.741% | 6.444% | 4.016% | 47.14% | 23.88% | 7.746% | 10.54% | 0% | 6.736% | 1.425% |  | 3.822% | 18.02% | 2.814% | 1.648% | 18.63% | 9.775% | 3.642%
| PDM3 |  | -17.71% | 1454% | -10.42% | 723.2% | 39.58% | 6.246% | 27.08% | 87.49% | -8.336% | -16.67% | -25% | -12.5% | -4.17% | 3.169% |  | 38.75% | 2.08% | 28787% | 15566% | -8.336% | 981.2% | 329.7%
| Z-Score (Absolute) |  | 0.67 | 46.86 | 0.44 | 23.25 | 1.18 | 0.1 | 0.77 | NA | 0.37 | 0.64 | 0.91 | 0.51 | 0.24 | 0 | 16.85 | 0.49 | 0.04 | 929.52 | 502.58 | 0.37 | 31.58 | 10.54
Results from laboratory 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 238b (Certified Value 55 ppm) | [25 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 57.377 | 81.7259 | 52.1 | 65.46 | 47.14 | 68.9 | 56 | 53.7 | 53.2 | 57.4 | 59.6 | 54.3901 | 58.96 | 57.18 | 54.64 | 37.9123 | 0.2365 | [7 empty] | 1
| 1 | 2 | 57.619 | 81.1638 | 53.2 | 65.89 | 49.96 | 72.3 | 55 | 55.5 | 52.8 | 55 | 58.1 | 54.6214 | 52.65 | 52.16 | 51.25 | 35.3058 | 0.2395 | [7 empty] | 31
| 1 | 3 | 57.098 | 81.33 | 51 | 68.88 | 43.13 | 68.9 | 56 | 53 | 58.1 | 55.9 | 56.9 | 55.8137 | 58.54 | 55.28 | 51.58 | 35.2716 | 0.2445 | [7 empty] | 16
| 1 | 4 | 56.168 | 81.656 | 49 | 67.87 | 46.56 | 71.5 | 57 | 54.3 | 57.3 | 55.1 | 59.3 | 53.5921 | 57.53 | 52.7 | 50.4 | 40.7752 | 0.2188 | [7 empty] | 54.6957
| 1 | 5 | 57.958 | 81.0621 | 54.5 | 68.94 | 46.88 | 68.5 | 55 | 55.8 | 54.2 | 57.6 | 57.9 | 54.7833 | 54.62 | 53.55 | 52.26 | 41.4078 | 0.229 | [7 empty] | 81
| 1 | 6 | 56.182 | 81.425 | 54.9 | 67.71 | 53.32 | 72.4 | 58 | 56.7 | 52.9 | 54.5 | 57.6 | 54.1162 | 62.5 | 56.08 | 49.76 | 36.0462 | 0.2271
| Mean |  | 57.067 | 81.3938 | 52.45 | 67.4583 | 47.8317 | 70.4167 | 56.1667 | 54.8333 | 54.75 | 55.9167 | 58.2333 | 54.5528 | 57.4667 | 54.4917 | 51.6483 | 37.7865 | 0.23257
| Median |  | 57.2375 | 81.3775 | 52.65 | 67.79 | 47.01 | 70.2 | 56 | 54.9 | 53.7 | 55.5 | 58 | 54.5058 | 58.035 | 54.415 | 51.415 | 36.9793 | 0.23275
| Std Dev. |  | 0.7466 | 0.26342 | 2.2314 | 1.4766 | 3.4571 | 1.84 | 1.169 | 1.3995 | 2.3518 | 1.3075 | 1.0309 | 0.74669 | 3.4648 | 1.9952 | 1.7095 | 2.7411 | 0.0093603
| Rel.Std.Dev. |  | 1.308% | 0.3236% | 4.254% | 2.189% | 7.228% | 2.613% | 2.081% | 2.552% | 4.296% | 2.338% | 1.77% | 1.369% | 6.029% | 3.662% | 3.31% | 7.254% | 4.025%
| PDM3 |  | 4.335% | 48.81% | -4.106% | 23.33% | -12.55% | 28.74% | 2.689% | 0.2516% | 0.09927% | 2.232% | 6.468% | -0.2613% | 5.066% | -0.373% | -5.572% | -30.92% | -99.57%
| Z-Score (Absolute) |  | 0.57 | 6.8 | 0.61 | 3.23 | 1.79 | 3.99 | 0.34 | 0 | 0.02 | 0.28 | 0.87 | 0.07 | 0.67 | 0.09 | 0.82 | 4.37 | 13.98
| Analytical results for Co in OREAS 238b (Certified Value 15.2 ppm) | [25 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 2
| 1 | 1 | 14.88 | 16.308 | 15.9 | 16.7353 | 14.9 | 14.8 | 15.1 | 13.7 | 15.8 | 14.7 | 15.2 | 15.1 | 14.9 | 16.2252 | 14 | 13.92 | 14.9 | 22.74 | 18.8 | 14.7 | 16.15 | 14.1206 | 15.3225 |  | 1
| 1 | 2 | 14.98 | 16.443 | 15.3 | 16.8469 | 15 | 14.8 | 15 | 13.6 | 16 | 15.2 | 15.1 | 15.4 | 14.9 | 16.2732 | 13 | 14.41 | 13.6 | 19.5 | 18.7 | 14.6 | 16.15 | 13.6613 | 15.1195 |  | 32
| 1 | 3 | 15.19 | 16.83 | 16 | 16.8444 | 14.9 | 15.3 | 14.7 | 13.6 | 15.6 | 15.6 | 15.2 | 14.8 | 16.4 | 15.486 | 14 | 14.62 | 15.5 | 16.24 | 18.7 | 14.8 | 15.39 | 13.1097 | 15.3497 |  | 16
| 1 | 4 | 15.09 | 16.578 | 15.5 | 16.9323 | 15.1 | 14.5 | 15 | 14.1 | 15.7 | 15.6 | 14.8 | 15.4 | 16.1 | 16.1659 | 13 | 13.25 | 15 | 16.88 | 18.2 | 14.4 | 16.34 | 13.4763 | 15.2172 |  | 15.2363
| 1 | 5 | 14.95 | 16.047 | 16.3 | 15.5272 | 15 | 15.5 | 14.7 | 13.7 | 16.1 | 15.2 | 15 | 15 | 16.3 | 15.4755 | 13 | 16.67 | 14.6 | 22.2 | 19.3 | 14.2 | 16.245 | 13.9349 | 14.9921 |  | 82
| 1 | 6 | 14.68 | 16.722 | 16.4 | 16.9548 | 15.2 | 15.5 | 15.6 | 14 | 16.4 | 15.4 | 14.8 | 14.2 | 15.8 | 15.6734 | 13 | 14.05 | 15.1 | 23.69 | 18.5 | 14.6 | 16.15 | 14.0249 | 15.4876
| Mean |  | 14.9617 | 16.488 | 15.9 | 16.6402 | 15.0167 | 15.0667 | 15.0167 | 13.7833 | 15.9333 | 15.2833 | 15.0167 | 14.9833 | 15.7333 | 15.8832 | 13.3333 | 14.4867 | 14.7833 | 20.2083 | 18.7 | 14.55 | 16.0708 | 13.7213 | 15.2481
| Median |  | 14.965 | 16.5105 | 15.95 | 16.8457 | 15 | 15.05 | 15 | 13.7 | 15.9 | 15.3 | 15.05 | 15.05 | 15.95 | 15.9196 | 13 | 14.23 | 14.95 | 20.85 | 18.7 | 14.6 | 16.15 | 13.7981 | 15.2699
| Std Dev. |  | 0.17612 | 0.28591 | 0.43359 | 0.55073 | 0.1169 | 0.42269 | 0.33116 | 0.2137 | 0.29439 | 0.33714 | 0.18348 | 0.44907 | 0.67725 | 0.37868 | 0.5164 | 1.1689 | 0.64936 | 3.1567 | 0.36332 | 0.21679 | 0.34209 | 0.38338 | 0.17675
| Rel.Std.Dev. |  | 1.177% | 1.734% | 2.727% | 3.31% | 0.7785% | 2.805% | 2.205% | 1.55% | 1.848% | 2.206% | 1.222% | 2.997% | 4.305% | 2.384% | 3.873% | 8.069% | 4.393% | 15.62% | 1.943% | 1.49% | 2.129% | 2.794% | 1.159%
| PDM3 |  | -1.802% | 8.215% | 4.356% | 9.214% | -1.441% | -1.113% | -1.441% | -9.536% | 4.575% | 0.3089% | -1.441% | -1.66% | 3.262% | 4.246% | -12.49% | -4.92% | -2.973% | 32.63% | 22.73% | -4.504% | 5.478% | -9.943% | 0.07771%
| Z-Score (Absolute) |  | 0.21 | 1.4 | 0.78 | 1.56 | 0.15 | 0.1 | 0.15 | 1.45 | 0.82 | 0.13 | 0.15 | 0.18 | 0.61 | 0.77 | NA | 0.71 | 0.39 | 5.33 | 3.74 | 0.64 | 0.96 | 1.52 | 0.1
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 238b (Certified Value 106 ppm) | [25 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 112.3 | 104.4525 | 96 | 115.46 | 99 | 111 | 108 | 107 | 106 | 110 | 109 | 104 | 106.3 | 101.273 | 107 | 119.3 | 100 | 107.33 | 109 | 102.8 | 93.3134 | 110.4638 |  |  | 1
| 1 | 2 | 112 | 106.3715 | 98 | 114.3967 | 100 | 108 | 111 | 108 | 106 | 110 | 107 | 105 | 106.4 | 108.956 | 104 | 130.1 | 96 | 100.19 | 107 | 100.8 | 94.1658 | 106.6029 |  |  | 33
| 1 | 3 | 112.1 | 109.3165 | 98 | 115.1733 | 99 | 109 | 111 | 106 | 104 | 112 | 109 | 106 | 107.4 | 106.7008 | 107 | 122.4 | 101 | 100.02 | 107 | 98 | 95.6109 | 110.7031 |  |  | 16
| 1 | 4 | 112.6 | 105.982 | 92 | 113.61 | 99 | 107 | 111 | 109 | 105 | 112 | 108 | 106 | 109.4 | 103.067 | 103 | 116 | 100 | 100.89 | 105 | 100.9 | 91.2218 | 107.5561 |  |  | 106.0053
| 1 | 5 | 110.7 | 105.7445 | 96 | 113.0233 | 100 | 110 | 110 | 108 | 106 | 113 | 109 | 104 | 104.8 | 101.9756 | 106 | 120.8 | 96 | 102.16 | 105 | 101.7 | 91.9371 | 108.7505 |  |  | 83
| 1 | 6 | 108.1 | 106.533 | 94 | 115.86 | 100 | 109 | 112 | 115 | 106 | 112 | 108 | 105 | 103.6 | 105.9172 | 105 | 121.3 | 99 | 107.16 | 107 | 100.9 | 91.7639 | 110.7518
| Mean |  | 111.3 | 106.4 | 95.6667 | 114.5872 | 99.5 | 109 | 110.5 | 108.8333 | 105.5 | 111.5 | 108.3333 | 105 | 106.3167 | 104.6483 | 105.3333 | 121.65 | 98.6667 | 102.9583 | 106.6667 | 100.85 | 93.0022 | 109.138
| Median |  | 112.05 | 106.1767 | 96 | 114.785 | 99.5 | 109 | 111 | 108 | 106 | 112 | 108.5 | 105 | 106.35 | 104.4921 | 105.5 | 121.05 | 99.5 | 101.525 | 107 | 100.9 | 92.6252 | 109.6072
| Std Dev. |  | 1.6982 | 1.6076 | 2.3381 | 1.1102 | 0.54772 | 1.4142 | 1.3784 | 3.1885 | 0.83666 | 1.2247 | 0.8165 | 0.89443 | 2.0183 | 3.0138 | 1.633 | 4.6958 | 2.1602 | 3.4052 | 1.5055 | 1.5909 | 1.6789 | 1.7827
| Rel.Std.Dev. |  | 1.526% | 1.511% | 2.444% | 0.9688% | 0.5505% | 1.297% | 1.247% | 2.93% | 0.793% | 1.098% | 0.7537% | 0.8518% | 1.898% | 2.88% | 1.55% | 3.86% | 2.189% | 3.307% | 1.411% | 1.578% | 1.805% | 1.633%
| PDM3 |  | 4.995% | 0.3724% | -9.753% | 8.096% | -6.137% | 2.825% | 4.24% | 2.668% | -0.4766% | 5.183% | 2.196% | -0.9483% | 0.2938% | -1.28% | -0.6339% | 14.76% | -6.923% | -2.874% | 0.6239% | -4.863% | -12.27% | 2.955%
| Z-Score (Absolute) |  | 1.08 | 0.01 | 2.33 | 1.8 | 1.5 | 0.58 | 0.9 | 0.54 | 0.19 | 1.12 | 0.43 | 0.3 | 0.01 | 0.37 | 0.22 | 3.34 | 1.68 | 0.74 | 0.07 | 1.2 | 2.91 | 0.61
| Analytical results for Cs in OREAS 238b (Certified Value 6.85 ppm) | [25 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.20 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 6.76 | 5.8811 | 6.86 | 6.81 | 6.55 | 7.67 | 6.7 | 7.02 | 6.92 | 6.85 | 7.05 | 7.221 | 6.63 | 7.08 | 6.8076 | [9 empty] | 1
| 1 | 2 | 6.8 | 6.3021 | 6.99 | 6.97 | 6.41 | 7.94 | 6.6 | 7.11 | 6.91 | 6.88 | 7.23 | 7.808 | 6.27 | 6.83 | 6.6593 | [9 empty] | 34
| 1 | 3 | 6.88 | 6.138 | 6.79 | 6.84 | 6.66 | 7.53 | 6.6 | 6.99 | 7.23 | 6.73 | 6.92 | 7.003 | 6.86 | 6.85 | 6.3875 | [9 empty] | 16
| 1 | 4 | 6.76 | 6.0855 | 6.33 | 6.89 | 6.42 | 7.29 | 6.8 | 6.76 | 7.24 | 6.73 | 7.15 | 6.778 | 6.36 | 6.69 | 7.3645 | [9 empty] | 6.8488
| 1 | 5 | 6.81 | 5.8587 | 6.85 | 6.91 | 6.73 | 7.66 | 6.7 | 7.19 | 6.98 | 6.97 | 7 | 7.503 | 6.21 | 6.45 | 7.4778 | [9 empty] | 84
| 1 | 6 | 6.8 | 6.3412 | 6.79 | 6.98 | 6.6 | 7.96 | 6.8 | 7.17 | 6.98 | 6.76 | 6.93 | 6.984 | 6.82 | 6.71 | 6.5431
| Mean |  | 6.8017 | 6.1011 | 6.7683 | 6.9 | 6.5617 | 7.675 | 6.7 | 7.04 | 7.0433 | 6.82 | 7.0467 | 7.2162 | 6.525 | 6.7683 | 6.8733
| Median |  | 6.8 | 6.1117 | 6.82 | 6.9 | 6.575 | 7.665 | 6.7 | 7.065 | 6.98 | 6.805 | 7.025 | 7.112 | 6.495 | 6.77 | 6.7335
| Std Dev. |  | 0.044008 | 0.20339 | 0.22684 | 0.068118 | 0.12859 | 0.25336 | 0.089443 | 0.15849 | 0.15135 | 0.096747 | 0.1234 | 0.3805 | 0.28346 | 0.20904 | 0.44763
| Rel.Std.Dev. |  | 0.647% | 3.334% | 3.351% | 0.9872% | 1.96% | 3.301% | 1.335% | 2.251% | 2.149% | 1.419% | 1.751% | 5.273% | 4.344% | 3.088% | 6.513%
| PDM3 |  | -0.6875% | -10.92% | -1.174% | 0.7483% | -4.192% | 12.06% | -2.172% | 2.792% | 2.841% | -0.4198% | 2.89% | 5.365% | -4.727% | -1.174% | 0.3584%
| Z-Score (Absolute) |  | 0.06 | 2.2 | 0.16 | 0.25 | 0.79 | 2.62 | 0.37 | 0.67 | 0.68 | 0 | 0.69 | 1.21 | 0.9 | 0.16 | 0.16
| Analytical results for Cu in OREAS 238b (Certified Value 29.6 ppm) | [25 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 2
| 1 | 1 | 28.99 | 26.771 | 30.5 | 28.3733 | 30 | 26.6 | 29 | 30.3 | 30.9 | 29.3 | 30.5 | 31.3 | 30.21 | 31.1708 | 32 | 27.93 | 27.5 | 28.63 | 31.9 | 30.4 | 30.457 | 29 | 1 | 30.0972 | 1
| 1 | 2 | 30.17 | 27.761 | 29.4 | 28.5033 | 32 | 27.1 | 29 | 27.9 | 30.3 | 29.2 | 29.1 | 31.8 | 29.19 | 32.1724 | 31 | 28.55 | 26.3 | 27.69 | 32 | 30 | 30.8845 | 30 | 1 | 29.5365 | 35
| 1 | 3 | 29.76 | 28.427 | 30.1 | 28.42 | 31 | 27.1 | 30 | 27.8 | 29.8 | 29.7 | 29 | 30.4 | 29.96 | 33.084 | 32 | 29.56 | 27.5 | 26.99 | 32.1 | 30.6 | 29.374 | 30 | 2 | 30.45 | 16
| 1 | 4 | 29.2 | 27.635 | 28.3 | 28.5367 | 32 | 26.9 | 29 | 28.6 | 29.9 | 30.2 | 29.1 | 31.3 | 30.83 | 31.7783 | 30 | 30.79 | 27.8 | 27.85 | 32.1 | 29.6 | 30.609 | 30 | 1 | 29.8429 | 29.5904
| 1 | 5 | 29.02 | 27.149 | 30.8 | 28.7 | 31 | 28.1 | 28 | 27.8 | 30.7 | 29.1 | 29.3 | 31 | 28.84 | 31.3384 | 31 | 26.73 | 26.4 | 27.91 | 31.5 | 29.3 | 30.343 | 32 | 1 | 29.8266 | 85
| 1 | 6 | 29.17 | 27.797 | 29.9 | 28.74 | 31 | 28.1 | 27 | 30.4 | 31.2 | 29.3 | 29 | 30.1 | 28.57 | 31.2844 | 31 | 31.45 | 27.3 | 29.65 | 31.7 | 30.1 | 30.1055 | 31 | 1 | 30.5962
| Mean |  | 29.385 | 27.59 | 29.8333 | 28.5456 | 31.1667 | 27.3167 | 28.6667 | 28.8 | 30.4667 | 29.4667 | 29.3333 | 30.9833 | 29.6 | 31.8047 | 31.1667 | 29.1683 | 27.1333 | 28.12 | 31.8833 | 30 | 30.2955 | 30.3333 | 1.1667 | 30.0582
| Median |  | 29.185 | 27.698 | 30 | 28.52 | 31 | 27.1 | 29 | 28.25 | 30.5 | 29.3 | 29.1 | 31.15 | 29.575 | 31.5583 | 31 | 29.055 | 27.4 | 27.88 | 31.95 | 30.05 | 30.4 | 30 | 1 | 29.9701
| Std Dev. |  | 0.47467 | 0.5727 | 0.89368 | 0.14761 | 0.75277 | 0.63377 | 1.0328 | 1.2377 | 0.56095 | 0.41312 | 0.58195 | 0.63061 | 0.87421 | 0.72969 | 0.75277 | 1.7805 | 0.62823 | 0.91431 | 0.24014 | 0.4858 | 0.52131 | 1.0328 | 0.40825 | 0.40414
| Rel.Std.Dev. |  | 1.615% | 2.076% | 2.996% | 0.5171% | 2.415% | 2.32% | 3.603% | 4.298% | 1.841% | 1.402% | 1.984% | 2.035% | 2.953% | 2.294% | 2.415% | 6.104% | 2.315% | 3.251% | 0.7532% | 1.619% | 1.721% | 3.405% | 34.99% | 1.345%
| PDM3 |  | -0.6943% | -6.76% | 0.8208% | -3.531% | 5.327% | -7.684% | -3.122% | -2.671% | 2.961% | -0.4183% | -0.8689% | 4.707% | 0.03228% | 7.483% | 5.327% | -1.427% | -8.304% | -4.969% | 7.749% | 1.384% | 2.383% | 2.511% | -96.06% | 1.581%
| Z-Score (Absolute) |  | 0.11 | 1.46 | 0.22 | 0.74 | 1.22 | 1.66 | 0.65 | 0.55 | 0.7 | 0.05 | 0.15 | 1.09 | 0.05 | 1.7 | 1.22 | 0.27 | 1.8 | 1.06 | 1.76 | 0.35 | 0.57 | 0.6 | 21.25 | 0.39
| Analytical results for Dy in OREAS 238b (Indicative Value 2.05 ppm) | [25 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 2.276; 1.06: 2; 1.07: 2.146; 1.08: 1.8433 |  |  |  | 1
| 1 | 2 | 1.01: 2.2757; 1.06: 2.1; 1.07: 1.957; 1.08: 1.6919 |  |  |  | 9
| 1 | 3 | 1.01: 2.3261; 1.06: 2.2; 1.07: 2.072; 1.08: 1.6612 |  |  |  | 16
| 1 | 4 | 1.01: 2.335; 1.06: 1.8; 1.07: 2.054; 1.08: 1.9407 |  |  |  | 2.0462
| 1 | 5 | 1.01: 2.3333; 1.06: 2; 1.07: 1.98; 1.08: 1.9664 |  |  |  | 15
| 1 | 6 | 1.01: 2.3407; 1.06: 1.9; 1.07: 2.194; 1.08: 1.7161
| Mean |  | 1.01: 2.3145; 1.06: 2; 1.07: 2.0672; 1.08: 1.8033
| Median |  | 1.01: 2.3297; 1.06: 2; 1.07: 2.063; 1.08: 1.7797
| Std Dev. |  | 1.01: 0.030239; 1.06: 0.14142; 1.07: 0.091905; 1.08: 0.13213
| Rel.Std.Dev. |  | 1.01: 1.307%; 1.06: 7.071%; 1.07: 4.446%; 1.08: 7.327%
| PDM3 |  | 1.01: 13.11%; 1.06: -2.259%; 1.07: 1.024%; 1.08: -11.87%
| Z-Score (Absolute) |  | 1.01: 1.44; 1.06: 0.17; 1.07: 0.17; 1.08: 1.18
| Analytical results for Er in OREAS 238b (Indicative Value 0.9 ppm) | [25 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.95468; 1.06: 0.9; 1.07: 0.964; 1.08: 0.7824 |  |  |  | 1
| 1 | 2 | 1.01: 1.0093; 1.06: 0.9; 1.07: 0.875; 1.08: 0.7562 |  |  |  | 10
| 1 | 3 | 1.01: 1.0194; 1.06: 0.8; 1.07: 0.938; 1.08: 0.7433 |  |  |  | 16
| 1 | 4 | 1.01: 0.99181; 1.06: 0.8; 1.07: 0.92; 1.08: 0.8477 |  |  |  | 0.90444
| 1 | 5 | 1.01: 1.0714; 1.06: 1; 1.07: 0.896; 1.08: 0.8803 |  |  |  | 16
| 1 | 6 | 1.01: 1.0091; 1.06: 0.9; 1.07: 0.99; 1.08: 0.7579
| Mean |  | 1.01: 1.0093; 1.06: 0.88333; 1.07: 0.9305; 1.08: 0.79463
| Median |  | 1.01: 1.0092; 1.06: 0.9; 1.07: 0.929; 1.08: 0.77015
| Std Dev. |  | 1.01: 0.038087; 1.06: 0.075277; 1.07: 0.04266; 1.08: 0.056153
| Rel.Std.Dev. |  | 1.01: 3.774%; 1.06: 8.522%; 1.07: 4.585%; 1.08: 7.067%
| PDM3 |  | 1.01: 11.59%; 1.06: -2.334%; 1.07: 2.881%; 1.08: -12.14%
| Z-Score (Absolute) |  | 1.01: 1.1; 1.06: 0.25; 1.07: 0.25; 1.08: 1.2
| Analytical results for Eu in OREAS 238b (Indicative Value 0.65 ppm) | [25 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.81492; 1.06: 0.6; 1.07: 0.678; 1.08: 0.5246 |  |  |  | 1
| 1 | 2 | 1.01: 0.81155; 1.06: 0.6; 1.07: 0.612; 1.08: 0.4972 |  |  |  | 11
| 1 | 3 | 1.01: 0.81632; 1.06: 0.6; 1.07: 0.666; 1.08: 0.5025 |  |  |  | 16
| 1 | 4 | 1.01: 0.84408; 1.06: 0.6; 1.07: 0.653; 1.08: 0.5718 |  |  |  | 0.65371
| 1 | 5 | 1.01: 0.84918; 1.06: 0.6; 1.07: 0.628; 1.08: 0.589 |  |  |  | 17
| 1 | 6 | 1.01: 0.81759; 1.06: 0.6; 1.07: 0.704; 1.08: 0.5092
| Mean |  | 1.01: 0.82561; 1.06: 0.6; 1.07: 0.65683; 1.08: 0.53238
| Median |  | 1.01: 0.81696; 1.06: 0.6; 1.07: 0.6595; 1.08: 0.5169
| Std Dev. |  | 1.01: 0.016487; 1.06: 0; 1.07: 0.033505; 1.08: 0.038699
| Rel.Std.Dev. |  | 1.01: 1.997%; 1.06: 0%; 1.07: 5.101%; 1.08: 7.269%
| PDM3 |  | 1.01: 26.3%; 1.06: -8.216%; 1.07: 0.4784%; 1.08: -18.56%
| Z-Score (Absolute) |  | 1.01: 2.14; 1.06: 0.31; 1.07: 0.31; 1.08: 1.04
| Analytical results for Fe in OREAS 238b (Certified Value 3.51 wt.%) | [25 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.397 | 3.6403 | 3.53 | 3.63 | 3.49 | 3.44 | 3.73 | 3.54 | 3.42 | 3.5 | 3.5 | 3.34 | 3.56 | 3.666 | 3.567 | 3.27 | 3.4884 | 3.61 | 3.61 | 3.3708 | 4.1406 |  |  |  | 1
| 1 | 2 | 3.483 | 3.5939 | 3.42 | 3.6367 | 3.52 | 3.5 | 3.84 | 3.55 | 3.45 | 3.53 | 3.45 | 3.38 | 3.59 | 3.651 | 3.483 | 3.15 | 3.5177 | 3.61 | 3.52 | 3.2946 | 4.202 |  |  |  | 0
| 1 | 3 | 3.463 | 3.6878 | 3.35 | 3.6733 | 3.51 | 3.53 | 3.83 | 3.52 | 3.41 | 3.57 | 3.49 | 3.4 | 3.6 | 3.6623 | 3.56 | 3.31 | 3.5038 | 3.62 | 3.55 | 3.212 | 4.2533 |  |  |  | 16
| 1 | 4 | 3.414 | 3.6089 | 3.34 | 3.64 | 3.52 | 3.46 | 3.85 | 3.56 | 3.44 | 3.56 | 3.49 | 3.42 | 3.67 | 3.6599 | 3.385 | 3.29 | 3.4907 | 3.61 | 3.48 | 3.3006 | 4.0571 |  |  |  | 3.5144
| 1 | 5 | 3.42 | 3.5691 | 3.45 | 3.64 | 3.49 | 3.56 | 3.84 | 3.53 | 3.47 | 3.62 | 3.51 | 3.38 | 3.52 | 3.6368 | 3.462 | 3.15 | 3.5309 | 3.6 | 3.45 | 3.3093 | 4.1172 |  |  |  | 86
| 1 | 6 | 3.355 | 3.6033 | 3.46 | 3.67 | 3.53 | 3.58 | 3.86 | 3.59 | 3.47 | 3.58 | 3.48 | 3.39 | 3.49 | 3.6643 | 3.497 | 3.25 | 3.5765 | 3.61 | 3.53 | 3.2881 | 4.0946
| Mean |  | 3.422 | 3.6172 | 3.425 | 3.6483 | 3.51 | 3.5117 | 3.825 | 3.5483 | 3.4433 | 3.56 | 3.4867 | 3.385 | 3.5717 | 3.6567 | 3.4923 | 3.2367 | 3.518 | 3.61 | 3.5233 | 3.2959 | 4.1441
| Median |  | 3.417 | 3.6061 | 3.435 | 3.64 | 3.515 | 3.515 | 3.84 | 3.545 | 3.445 | 3.565 | 3.49 | 3.385 | 3.575 | 3.6611 | 3.49 | 3.26 | 3.5108 | 3.61 | 3.525 | 3.2976 | 4.1289
| Std Dev. |  | 0.046009 | 0.04157 | 0.071764 | 0.018469 | 0.016733 | 0.055287 | 0.047645 | 0.024833 | 0.025033 | 0.041473 | 0.020656 | 0.026646 | 0.063692 | 0.011092 | 0.067384 | 0.070048 | 0.032921 | 0.0063246 | 0.055737 | 0.05082 | 0.072156
| Rel.Std.Dev. |  | 1.344% | 1.149% | 2.095% | 0.5062% | 0.4767% | 1.574% | 1.246% | 0.6998% | 0.727% | 1.165% | 0.5924% | 0.7872% | 1.783% | 0.3033% | 1.929% | 2.164% | 0.9358% | 0.1752% | 1.582% | 1.542% | 1.741%
| PDM3 |  | -2.628% | 2.927% | -2.543% | 3.812% | -0.1244% | -0.07697% | 8.839% | 0.9664% | -2.021% | 1.298% | -0.7883% | -3.681% | 1.63% | 4.051% | -0.6271% | -7.902% | 0.1037% | 2.721% | 0.255% | -6.217% | 17.92%
| Z-Score (Absolute) |  | 0.7 | 0.73 | 0.68 | 0.96 | 0.06 | 0.05 | 2.25 | 0.22 | 0.55 | 0.31 | 0.23 | 0.98 | 0.39 | 1.02 | 0.19 | 2.06 | 0 | 0.67 | 0.04 | 1.63 | 4.59
| Analytical results for Ga in OREAS 238b (Certified Value 7.93 ppm) | [25 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.21 | 1.25 | 1.27 | 1.30 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 8.3 | 7.1502 | 7.16 | 8.68 | 8.1 | 9 | 7.3 | 9 | 7.94 | 7.9 | 8.03 | 7.82 | 8.1 | 8.6912 | 8 | 11.6 | 9.09 | 6.638 | 7.7701 | [5 empty] | 1
| 1 | 2 | 8.3 | 7.3759 | 7.15 | 9.0613 | 8.1 | 10 | 7.4 | 9 | 8.06 | 7.87 | 7.95 | 7.93 | 8.4 | 8.4436 | 7.6 | 11.7 | 8.63 | 7.1117 | 7.824 | [5 empty] | 39
| 1 | 3 | 8.3 | 7.1752 | 7.08 | 8.7813 | 8.1 | 9 | 7.1 | 9 | 7.77 | 8.2 | 7.87 | 7.65 | 8.4 | 8.6775 | 7.7 | 11.4 | 8.78 | 6.8309 | 7.6058 | [5 empty] | 16
| 1 | 4 | 8.2 | 6.9919 | 6.89 | 8.8427 | 8.2 | 9 | 7.5 | 9 | 7.75 | 8.53 | 7.8 | 7.94 | 8.4 | 8.3752 | 7.5 | 10.5 | 8.72 | 7.8074 | 7.4362 | [5 empty] | 7.9265
| 1 | 5 | 8.3 | 7.0964 | 7.03 | 8.0373 | 8.2 | 9 | 7.3 | 9 | 8.13 | 7.84 | 8.07 | 7.81 | 8.6 | 8.582 | 7.6 | 10.9 | 8.37 | 7.9127 | 7.6738 | [5 empty] | 87
| 1 | 6 | 8.1 | 7.1358 | 6.93 | 9.088 | 8.2 | 8 | 7.7 | 9 | 8.14 | 8.11 | 7.82 | 7.64 | 8.4 | 8.2433 | 8.4 | 11.4 | 8.66 | 6.9036 | 7.6695
| Mean |  | 8.25 | 7.1542 | 7.04 | 8.7484 | 8.15 | 9 | 7.3833 | 9 | 7.965 | 8.075 | 7.9233 | 7.7983 | 8.3833 | 8.5021 | 7.8 | 11.25 | 8.7083 | 7.2007 | 7.6632
| Median |  | 8.3 | 7.143 | 7.055 | 8.812 | 8.15 | 9 | 7.35 | 9 | 8 | 8.005 | 7.91 | 7.815 | 8.4 | 8.5128 | 7.65 | 11.4 | 8.69 | 7.0076 | 7.6716
| Std Dev. |  | 0.083666 | 0.1262 | 0.11207 | 0.38302 | 0.054772 | 0.63246 | 0.20412 | 0 | 0.17421 | 0.26524 | 0.11165 | 0.13045 | 0.16021 | 0.17855 | 0.34059 | 0.45935 | 0.23404 | 0.53378 | 0.13585
| Rel.Std.Dev. |  | 1.014% | 1.764% | 1.592% | 4.378% | 0.6721% | 7.027% | 2.765% | 0% | 2.187% | 3.285% | 1.409% | 1.673% | 1.911% | 2.1% | 4.367% | 4.083% | 2.688% | 7.413% | 1.773%
| PDM3 |  | 4.081% | -9.743% | -11.18% | 10.37% | 2.819% | 13.54% | -6.853% | 13.54% | 0.4855% | 1.873% | -0.04021% | -1.617% | 5.763% | 7.262% | -1.596% | 41.93% | 9.863% | -9.157% | -3.322%
| Z-Score (Absolute) |  | 0.46 | 1.31 | 1.49 | 1.26 | 0.3 | NA | 0.94 | NA | 0 | 0.18 | 0.07 | 0.27 | 0.67 | 0.87 | 0.27 | 5.3 | 1.2 | 1.23 | 0.49
Results from laboratories 1.09 and 1.11 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 238b (Indicative Value 4.03 ppm) | [25 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 5.9572; 1.06: 3.4; 1.07: 3.91; 1.08: 3.2966 |  |  |  | 1
| 1 | 2 | 1.01: 5.8679; 1.06: 3; 1.07: 3.544; 1.08: 3.187 |  |  |  | 12
| 1 | 3 | 1.01: 5.9065; 1.06: 2.9; 1.07: 3.843; 1.08: 3.0698 |  |  |  | 16
| 1 | 4 | 1.01: 5.8001; 1.06: 3; 1.07: 3.761; 1.08: 3.5609 |  |  |  | 4.0345
| 1 | 5 | 1.01: 5.9761; 1.06: 3.4; 1.07: 3.603; 1.08: 3.6385 |  |  |  | 18
| 1 | 6 | 1.01: 5.866; 1.06: 3.1; 1.07: 4.057; 1.08: 3.1834
| Mean |  | 1.01: 5.8956; 1.06: 3.1333; 1.07: 3.7863; 1.08: 3.3227
| Median |  | 1.01: 5.8872; 1.06: 3.05; 1.07: 3.802; 1.08: 3.2418
| Std Dev. |  | 1.01: 0.065041; 1.06: 0.21602; 1.07: 0.19222; 1.08: 0.22756
| Rel.Std.Dev. |  | 1.01: 1.103%; 1.06: 6.894%; 1.07: 5.077%; 1.08: 6.849%
| PDM3 |  | 1.01: 46.13%; 1.06: -22.34%; 1.07: -6.151%; 1.08: -17.64%
| Z-Score (Absolute) |  | 1.01: 4.84; 1.06: 0.87; 1.07: 0.48; 1.08: 0.48
| Analytical results for Ge in OREAS 238b (Certified Value 0.11 ppm) | [25 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.06; 1.07: 1.09; 1.08: 1.11; 1.09: 1.12; 1.10: 1.13; 1.11: 1.14; 1.12: 1.15; 1.13: 1.21; 1.14: 1.27; 1.15: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*OES/MS; 1.09: AR*MS; 1.10: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*OES/MS; 1.15: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.2g; 1.08: 0.25g; 1.09: 0.5g; 1.10: 0.5g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.25g; 1.14: 0.5g; 1.15: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.11; 1.08: 1.7; 1.09: 0.19; 1.10: 0.11; 1.11: 0.07; 1.12: 0.12; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0663 |  |  |  | 1
| 1 | 2 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.11; 1.08: 1.7; 1.09: 0.19; 1.10: 0.11; 1.11: 0.06; 1.12: 0.11; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0628 |  |  |  | 5
| 1 | 3 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.12; 1.08: 1.7; 1.09: 0.19; 1.10: 0.11; 1.11: 0.07; 1.12: 0.12; 1.13: 0.2; 1.14: 0.12; 1.15: 0.0632 |  |  |  | 16
| 1 | 4 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.12; 1.08: 1.7; 1.09: 0.21; 1.10: 0.12; 1.11: 0.06; 1.12: 0.1; 1.13: 0.2; 1.14: 0.14; 1.15: 0.0776 |  |  |  | 0.10729
| 1 | 5 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.12; 1.08: 1.7; 1.09: 0.2; 1.10: 0.11; 1.11: 0.06; 1.12: 0.11; 1.13: 0.2; 1.14: 0.13; 1.15: 0.0755 |  |  |  | 88
| 1 | 6 | 1.01: < 0.05; 1.06: < 0.1; 1.07: 0.12; 1.08: 1.7; 1.09: 0.2; 1.10: 0.12; 1.11: 0.06; 1.12: 0.1; 1.13: 0.2; 1.14: 0.13; 1.15: 0.0733
| Mean |  | 1.07: 0.11667; 1.08: 1.7; 1.09: 0.19667; 1.10: 0.11333; 1.11: 0.063333; 1.12: 0.11; 1.13: 0.2; 1.14: 0.12667; 1.15: 0.069783
| Median |  | 1.07: 0.12; 1.08: 1.7; 1.09: 0.195; 1.10: 0.11; 1.11: 0.06; 1.12: 0.11; 1.13: 0.2; 1.14: 0.125; 1.15: 0.0698
| Std Dev. |  | 1.07: 0.005164; 1.08: 0; 1.09: 0.008165; 1.10: 0.005164; 1.11: 0.005164; 1.12: 0.0089443; 1.13: 3.0405e-17; 1.14: 0.008165; 1.15: 0.0064867
| Rel.Std.Dev. |  | 1.07: 4.426%; 1.08: 0%; 1.09: 4.152%; 1.10: 4.556%; 1.11: 8.154%; 1.12: 8.131%; 1.13: 1.52e-14%; 1.14: 6.446%; 1.15: 9.296%
| PDM3 |  | 1.07: 8.74%; 1.08: 1484%; 1.09: 83.3%; 1.10: 5.633%; 1.11: -40.97%; 1.12: 2.526%; 1.13: 86.41%; 1.14: 18.06%; 1.15: -34.96%
| Z-Score (Absolute) |  | 1.01: 1.43; 1.06: 1.01; 1.07: 0.11; 1.08: NA; 1.09: 1.45; 1.10: 0.06; 1.11: 0.78; 1.12: 0; 1.13: NA; 1.14: 0.28; 1.15: 0.67
Results from laboratories 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 238b (Certified Value 0.54 ppm) | [25 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | 1.32 | [9 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.51 | 0.7743 | < 0.1 | 0.65 | 0.5 | 0.69 | 0.5 | 0.55 | 0.51 | 0.52 | 0.58 | 0.56 | 0.33 | 0.3361 | 0.4213 | [9 empty] | 1
| 1 | 2 | 0.57 | 0.76815 | < 0.1 | 0.66 | 0.43 | 0.69 | 0.5 | 0.56 | 0.5 | 0.51 | 0.55 | 0.54 | 0.32 | 0.333 | 0.4405 | [9 empty] | 25
| 1 | 3 | 0.54 | 0.80309 | < 0.1 | 0.64 | 0.51 | 0.68 | 0.5 | 0.56 | 0.51 | 0.51 | 0.57 | 0.55 | 0.32 | 0.3249 | 0.4557 | [9 empty] | 16
| 1 | 4 | 0.54 | 0.78225 | < 0.1 | 0.64 | 0.49 | 0.64 | 0.5 | 0.55 | 0.52 | 0.51 | 0.55 | 0.5 | 0.3 | 0.3441 | 0.3733 | [9 empty] | 0.541
| 1 | 5 | 0.5 | 0.81641 | 0.2 | 0.66 | 0.4 | 0.66 | 0.5 | 0.55 | 0.5 | 0.53 | 0.57 | 0.5 | 0.28 | 0.3547 | 0.3609 | [9 empty] | 89
| 1 | 6 | 0.6 | 0.76652 | 0.2 | 0.66 | 0.47 | 0.67 | 0.5 | 0.56 | 0.53 | 0.51 | 0.57 | 0.52 | 0.29 | 0.3161 | 0.3581
| Mean |  | 0.54333 | 0.78512 | 0.2 | 0.65167 | 0.46667 | 0.67167 | 0.5 | 0.555 | 0.51167 | 0.515 | 0.565 | 0.52833 | 0.30667 | 0.33482 | 0.40163
| Median |  | 0.54 | 0.77827 | 0.2 | 0.655 | 0.48 | 0.675 | 0.5 | 0.555 | 0.51 | 0.51 | 0.57 | 0.53 | 0.31 | 0.33455 | 0.3973
| Std Dev. |  | 0.037238 | 0.020303 | 0 | 0.0098319 | 0.043205 | 0.019408 | 0 | 0.0054772 | 0.01169 | 0.0083666 | 0.012247 | 0.025626 | 0.019664 | 0.013676 | 0.042843
| Rel.Std.Dev. |  | 6.854% | 2.586% | 0% | 1.509% | 9.258% | 2.89% | 0% | 0.9869% | 2.285% | 1.625% | 2.168% | 4.85% | 6.412% | 4.085% | 10.67%
| PDM3 |  | 0.4319% | 45.12% | -63.03% | 20.46% | -13.74% | 24.15% | -7.578% | 2.588% | -5.421% | -4.805% | 4.437% | -2.341% | -43.31% | -38.11% | -25.76%
| Z-Score (Absolute) |  | 0.16 | 2.81 | NA | 1.35 | 0.67 | 1.57 | NA | 0.29 | 0.18 | 0.15 | 0.4 | 0 | 2.42 | 2.12 | 1.39
Results from laboratories 1.06 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 238b (Indicative Value 0.015 ppm) | [25 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 01g | 0.2g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.01 | 0.06 | < 0.05 | < 2 | 0.104 | 0.01 | 0.02 | 0.01 | 0.01 | 0.027 | < 1 | 0.02 | < 5 | 0.007 | 0.0156 | [9 empty] | 1
| 1 | 2 | 0.01 | 0.06 | < 0.05 | < 2 | 0.093 | 0.01 | 0.02 | 0.01 | 0.01 | 0.027 | < 1 | < 0.01 | < 5 | 0.01 | 0.0245 | [9 empty] | 6
| 1 | 3 | 0.01 | 0.07 | < 0.05 | < 2 | 0.092 | 0.01 | 0.02 | 0.01 | 0.01 | 0.025 | < 1 | 0.01 | < 5 | 0.011 | 0.0231 | [9 empty] | 16
| 1 | 4 | 0.01 | 0.05 | < 0.05 | < 2 | 0.105 | 0.01 | 0.02 | 0.01 | 0.01 | 0.026 | < 1 | < 0.01 | < 5 | 0.009 | 0.0274 | [9 empty] | 0.015272
| 1 | 5 | 0.01 | 0.06 | < 0.05 | < 2 | 0.094 | 0.02 | 0.02 | 0.01 | 0.01 | 0.028 | < 1 | 0.01 | < 5 | 0.009 | 0.0216 | [9 empty] | 19
| 1 | 6 | 0.01 | 0.04 | < 0.05 | < 2 | 0.102 | 0.02 | 0.02 | 0.02 | 0.01 | 0.022 | < 1 | 0.02 | < 5 | 0.012 | 0.0195
| Mean |  | 1.01: 0.01; 1.06: 0.056667; 1.09: 0.098333; 1.10: 0.013333; 1.11: 0.02; 1.12: 0.011667; 1.13: 0.01; 1.14: 0.025833; 1.16: 0.015; 1.19: 0.0096667; 1.21: 0.02195
| Median |  | 1.01: 0.01; 1.06: 0.06; 1.09: 0.098; 1.10: 0.01; 1.11: 0.02; 1.12: 0.01; 1.13: 0.01; 1.14: 0.0265; 1.16: 0.015; 1.19: 0.0095; 1.21: 0.02235
| Std Dev. |  | 1.01: 0; 1.06: 0.010328; 1.09: 0.0059554; 1.10: 0.005164; 1.11: 0; 1.12: 0.0040825; 1.13: 0; 1.14: 0.002137; 1.16: 0.0057735; 1.19: 0.0017512; 1.21: 0.0040982
| Rel.Std.Dev. |  | 1.01: 0%; 1.06: 18.23%; 1.09: 6.056%; 1.10: 38.73%; 1.11: 0%; 1.12: 34.99%; 1.13: 0%; 1.14: 8.272%; 1.16: 38.49%; 1.19: 18.12%; 1.21: 18.67%
| PDM3 |  | 1.01: -34.52%; 1.06: 271%; 1.09: 543.9%; 1.10: -12.7%; 1.11: 30.96%; 1.12: -23.61%; 1.13: -34.52%; 1.14: 69.15%; 1.16: -1.782%; 1.19: -36.7%; 1.21: 43.72%
| Z-Score (Absolute) |  | 0.67 | 1.96 | 0.17 | 55.19 | 4.31 | 0.49 | 0.11 | 0.58 | 0.67 | 0.22 | 26.98 | 0.58 | 139.83 | 0.69 | 0
| Analytical results for Ho in OREAS 238b (Indicative Value 0.34 ppm) | [25 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.37807; 1.06: 0.3; 1.07: 0.364; 1.08: 0.3037 |  |  |  | 1
| 1 | 2 | 1.01: 0.38969; 1.06: 0.3; 1.07: 0.331; 1.08: 0.2913 |  |  |  | 14
| 1 | 3 | 1.01: 0.4005; 1.06: 0.3; 1.07: 0.351; 1.08: 0.2808 |  |  |  | 16
| 1 | 4 | 1.01: 0.39755; 1.06: 0.3; 1.07: 0.345; 1.08: 0.3275 |  |  |  | 0.33768
| 1 | 5 | 1.01: 0.41619; 1.06: 0.3; 1.07: 0.336; 1.08: 0.3333 |  |  |  | 20
| 1 | 6 | 1.01: 0.39804; 1.06: 0.3; 1.07: 0.373; 1.08: 0.2877
| Mean |  | 1.01: 0.39668; 1.06: 0.3; 1.07: 0.35; 1.08: 0.30405
| Median |  | 1.01: 0.3978; 1.06: 0.3; 1.07: 0.348; 1.08: 0.2975
| Std Dev. |  | 1.01: 0.012594; 1.06: 0; 1.07: 0.016174; 1.08: 0.021799
| Rel.Std.Dev. |  | 1.01: 3.175%; 1.06: 0%; 1.07: 4.621%; 1.08: 7.17%
| PDM3 |  | 1.01: 17.47%; 1.06: -11.16%; 1.07: 3.648%; 1.08: -9.96%
| Z-Score (Absolute) |  | 1.01: 1.88; 1.06: 0.73; 1.07: 0.62; 1.08: 0.62
| Analytical results for In in OREAS 238b (Certified Value 0.029 ppm) | [25 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 0.03 | 0.03 | < 5 | < 0.05 | 0.017 | < 0.05 | 0.03 | 0.031 | 0.028 | 0.028 | 0.03 | 0.03 | 0.027 | 0.0349 | [10 empty] | 1
| 1 | 2 | 0.03 | 0.03 | < 5 | < 0.05 | 0.02 | < 0.05 | 0.03 | 0.03 | 0.028 | 0.028 | 0.03 | 0.02 | 0.028 | 0.0316 | [10 empty] | 27
| 1 | 3 | 0.03 | 0.03 | < 5 | < 0.05 | 0.019 | < 0.05 | 0.03 | 0.028 | 0.028 | 0.026 | 0.029 | 0.03 | 0.027 | 0.0234 | [10 empty] | 16
| 1 | 4 | 0.03 | 0.03 | < 5 | < 0.05 | 0.018 | < 0.05 | 0.03 | 0.029 | 0.029 | 0.025 | 0.031 | 0.03 | 0.027 | 0.0351 | [10 empty] | 0.029382
| 1 | 5 | 0.03 | 0.03 | < 5 | < 0.05 | 0.016 | < 0.05 | 0.03 | 0.032 | 0.027 | 0.027 | 0.033 | 0.03 | 0.027 | 0.0332 | [10 empty] | 90
| 1 | 6 | 0.03 | 0.04 | < 5 | < 0.05 | 0.019 | < 0.05 | 0.03 | 0.032 | 0.028 | 0.028 | 0.028 | 0.03 | 0.027 | 0.0287
| Mean |  | 1.01: 0.03; 1.06: 0.031667; 1.09: 0.018167; 1.11: 0.03; 1.12: 0.030333; 1.13: 0.028; 1.14: 0.027; 1.15: 0.030167; 1.16: 0.028333; 1.17: 0.027167; 1.19: 0.03115
| Median |  | 1.01: 0.03; 1.06: 0.03; 1.09: 0.0185; 1.11: 0.03; 1.12: 0.0305; 1.13: 0.028; 1.14: 0.0275; 1.15: 0.03; 1.16: 0.03; 1.17: 0.027; 1.19: 0.0324
| Std Dev. |  | 1.01: 0; 1.06: 0.0040825; 1.09: 0.001472; 1.11: 0; 1.12: 0.001633; 1.13: 0.00063246; 1.14: 0.0012649; 1.15: 0.0017224; 1.16: 0.0040825; 1.17: 0.00040825; 1.19: 0.0044752
| Rel.Std.Dev. |  | 1.01: 0%; 1.06: 12.89%; 1.09: 8.103%; 1.11: 0%; 1.12: 5.383%; 1.13: 2.259%; 1.14: 4.685%; 1.15: 5.71%; 1.16: 14.41%; 1.17: 1.503%; 1.19: 14.37%
| PDM3 |  | 1.01: 2.104%; 1.06: 7.777%; 1.09: -38.17%; 1.11: 2.104%; 1.12: 3.239%; 1.13: -4.702%; 1.14: -8.106%; 1.15: 2.672%; 1.16: -3.568%; 1.17: -7.539%; 1.19: 6.018%
| Z-Score (Absolute) |  | 0.28 | 0.85 | 836.54 | 1.41 | 3.72 | 1.41 | 0.28 | 0.39 | 0.39 | 0.73 | 0.34 | 0.28 | 0.68 | 0.67
| Analytical results for Ir in OREAS 238b (Indicative Value < 0.0015 ppm) | [25 empty] | Indicative
Ir | Round | Replicate | 1.01: Lab |  |  |  | 1
| No | No | 1.01: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS |  |  |  | 3
|  |  | 1.01: 0.5g |  |  |  | 3
| 1 | 1 | 1.01: < 0.0021 |  |  |  | 1
| 1 | 2 | 1.01: < 0.0015 |  |  |  | 3
| 1 | 3 | 1.01: < 0.0018 |  |  |  | 16
| 1 | 4 | 1.01: < 0.0016 |  |  |  | < 0.0015
| 1 | 5 | 1.01: < 0.0013 |  |  |  | 14
| 1 | 6 | 1.01: < 0.0015
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for K in OREAS 238b (Certified Value 0.87 wt.%) | [25 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.838 | 1.0178 | 0.77 | 0.888 | 0.909 | 0.89 | 0.977 | 0.84 | 0.84 | 0.9 | 0.9 | 0.84 | 0.87 | 0.85475 | 0.872 | 0.81 | 0.84557 | 0.87 | 0.9 | 0.9012 | 1.008 |  |  |  | 1
| 1 | 2 | 0.843 | 1.0262 | 0.79 | 0.902 | 0.881 | 0.86 | 0.963 | 0.84 | 0.84 | 0.91 | 0.89 | 0.85 | 0.88 | 0.87163 | 0.855 | 0.78 | 0.8518 | 0.88 | 0.88 | 0.8686 | 1.0077 |  |  |  | 0
| 1 | 3 | 0.85 | 1.0532 | 0.8 | 0.893 | 0.896 | 0.86 | 0.954 | 0.83 | 0.84 | 0.92 | 0.9 | 0.85 | 0.89 | 0.86257 | 0.872 | 0.81 | 0.84392 | 0.87 | 0.89 | 0.851 | 1.0077 |  |  |  | 16
| 1 | 4 | 0.849 | 1.0242 | 0.76 | 0.9065 | 0.884 | 0.86 | 0.961 | 0.84 | 0.83 | 0.92 | 0.9 | 0.86 | 0.9 | 0.84704 | 0.838 | 0.8 | 0.83873 | 0.87 | 0.86 | 0.8729 | 1.0093 |  |  |  | 0.87015
| 1 | 5 | 0.835 | 0.99935 | 0.77 | 0.901 | 0.907 | 0.87 | 0.951 | 0.83 | 0.85 | 0.94 | 0.9 | 0.85 | 0.86 | 0.84904 | 0.855 | 0.77 | 0.84725 | 0.87 | 0.85 | 0.882 | 1.0081 |  |  |  | 91
| 1 | 6 | 0.815 | 1.0098 | 0.77 | 0.91 | 0.883 | 0.88 | 0.967 | 0.85 | 0.85 | 0.91 | 0.89 | 0.84 | 0.85 | 0.85374 | 0.863 | 0.81 | 0.85795 | 0.87 | 0.87 | 0.878 | 1.0077
| Mean |  | 0.83833 | 1.0217 | 0.77667 | 0.90008 | 0.89333 | 0.87 | 0.96217 | 0.83833 | 0.84167 | 0.91667 | 0.89667 | 0.84833 | 0.875 | 0.85646 | 0.85917 | 0.79667 | 0.84754 | 0.87167 | 0.875 | 0.87562 | 1.0081
| Median |  | 0.8405 | 1.021 | 0.77 | 0.9015 | 0.89 | 0.865 | 0.962 | 0.84 | 0.84 | 0.915 | 0.9 | 0.85 | 0.875 | 0.85425 | 0.859 | 0.805 | 0.84641 | 0.87 | 0.875 | 0.87545 | 1.0078
| Std Dev. |  | 0.012863 | 0.018298 | 0.015055 | 0.0082487 | 0.012533 | 0.012649 | 0.009347 | 0.0075277 | 0.0075277 | 0.013663 | 0.005164 | 0.0075277 | 0.018708 | 0.0091829 | 0.012859 | 0.017512 | 0.0066525 | 0.0040825 | 0.018708 | 0.016514 | 0.00062102
| Rel.Std.Dev. |  | 1.534% | 1.791% | 1.938% | 0.9164% | 1.403% | 1.454% | 0.9715% | 0.8979% | 0.8944% | 1.49% | 0.5759% | 0.8874% | 2.138% | 1.072% | 1.497% | 2.198% | 0.7849% | 0.4684% | 2.138% | 1.886% | 0.0616%
| PDM3 |  | -3.657% | 17.42% | -10.74% | 3.44% | 2.664% | -0.01747% | 10.57% | -3.657% | -3.274% | 5.346% | 3.047% | -2.507% | 0.5571% | -1.573% | -1.262% | -8.445% | -2.599% | 0.1741% | 0.5571% | 0.628% | 15.85%
| Z-Score (Absolute) |  | 0.9 | 4.05 | 2.56 | 0.77 | 0.58 | 0.04 | 2.44 | 0.9 | 0.81 | 1.21 | 0.67 | 0.63 | 0.09 | 0.41 | 0.34 | 2.02 | 0.65 | 0 | 0.09 | 0.11 | 3.68
| Analytical results for La in OREAS 238b (Certified Value 27.1 ppm) | [25 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 28.154 | 35.1912 | 24.5 | 26.9633 | 33.88 | 25 | 33.9 | 30 | 26.7 | 25.6 | 28.1 | 29 | 26.4 | 26.8702 | 18 | 29.1 | 28.1 | 32.2 | 24.3 | 21.7054 |  |  |  |  | 1
| 1 | 2 | 28.906 | 35.1401 | 25.1 | 26.7967 | 34.14 | 21 | 35.1 | 29 | 27.6 | 25.5 | 27.4 | 29.4 | 27.1 | 27.1261 | 18 | 25.9 | 26.8 | 31.5 | 22 | 20.0728 |  |  |  |  | 28
| 1 | 3 | 28.661 | 35.7232 | 24.9 | 26.45 | 34.77 | 22 | 33.7 | 27 | 26.4 | 27.8 | 28.2 | 28.5 | 28 | 27.0607 | 18 | 29.04 | 28.6 | 31.5 | 22.9 | 20.0676 |  |  |  |  | 16
| 1 | 4 | 27.982 | 35.7334 | 23.5 | 26.2133 | 34.13 | 24 | 35.1 | 28 | 27 | 27 | 27.4 | 29.8 | 28.1 | 27.059 | 18 | 28.23 | 26.2 | 32.3 | 21.9 | 23.0927 |  |  |  |  | 27.0633
| 1 | 5 | 28.57 | 35.1503 | 25.9 | 26.4267 | 34.5 | 24 | 33.3 | 28 | 27.8 | 25.8 | 28.3 | 29.1 | 27 | 26.7505 | 18 | 26.9 | 26.3 | 32.9 | 20.7 | 23.4559 |  |  |  |  | 92
| 1 | 6 | 27.896 | 35.9789 | 25.8 | 27.4267 | 34.5 | 25 | 35.4 | 29 | 28.1 | 25.3 | 27.2 | 28.3 | 25.9 | 26.9762 | 18 | 30.74 | 27.2 | 31.3 | 21.3 | 20.3337
| Mean |  | 28.3615 | 35.4862 | 24.95 | 26.7128 | 34.32 | 23.5 | 34.4167 | 28.5 | 27.2667 | 26.1667 | 27.7667 | 29.0167 | 27.0833 | 26.9738 | 18 | 28.3183 | 27.2 | 31.95 | 22.1833 | 21.4547
| Median |  | 28.362 | 35.4572 | 25 | 26.6233 | 34.32 | 24 | 34.5 | 28.5 | 27.3 | 25.7 | 27.75 | 29.05 | 27.05 | 27.0176 | 18 | 28.635 | 27 | 31.85 | 21.95 | 21.0195
| Std Dev. |  | 0.40827 | 0.3687 | 0.88938 | 0.44268 | 0.32539 | 1.6432 | 0.88638 | 1.0488 | 0.66833 | 1.0013 | 0.48442 | 0.55648 | 0.86583 | 0.14044 | 0 | 1.7242 | 0.97365 | 0.61887 | 1.2719 | 1.5389
| Rel.Std.Dev. |  | 1.44% | 1.039% | 3.565% | 1.657% | 0.9481% | 6.992% | 2.575% | 3.68% | 2.451% | 3.827% | 1.745% | 1.918% | 3.197% | 0.5206% | 0% | 6.089% | 3.58% | 1.937% | 5.733% | 7.173%
| PDM3 |  | 4.797% | 31.12% | -7.809% | -1.295% | 26.81% | -13.17% | 27.17% | 5.309% | 0.7514% | -3.313% | 2.599% | 7.218% | 0.07397% | -0.3309% | -33.49% | 4.637% | 0.5051% | 18.06% | -18.03% | -20.72%
| Z-Score (Absolute) |  | 0.5 | 3.65 | 1.01 | 0.23 | 3.13 | 1.65 | 3.18 | 0.56 | 0.01 | 0.47 | 0.24 | 0.79 | 0.07 | 0.11 | 4.08 | 0.48 | 0.01 | 2.09 | 2.23 | 2.56
| Analytical results for Li in OREAS 238b (Certified Value 38.9 ppm) | [25 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [6 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | [6 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 38.23 | 43.6914 | 41.2 | 41.345 | 35.9 | 41 | 41.5 | 41.8 | 40.7 | 36.9 | 38.8 | 41.8 | 40.0925 | 38.29 | 36 | 37.6 | 36 | 38.6757 | [6 empty] | 1
| 1 | 2 | 37.63 | 44.0095 | 41.1 | 41.11 | 35 | 40 | 40.6 | 40.6 | 41.3 | 36.5 | 38.1 | 41.5 | 38.0494 | 40.82 | 35 | 35.5 | 34.3 | 37.7364 | [6 empty] | 0
| 1 | 3 | 38.13 | 43.726 | 41 | 41.37 | 36.7 | 40 | 39.5 | 38.8 | 40.4 | 38.4 | 38.6 | 39.1 | 40.088 | 38.98 | 37 | 35.6 | 33.7 | 39.0094 | [6 empty] | 16
| 1 | 4 | 38.13 | 43.9864 | 38.7 | 41.62 | 35.3 | 40 | 43 | 40 | 40.3 | 38.6 | 37.4 | 39.8 | 37.9192 | 37.27 | 36 | 35.6 | 34.7 | 36.2353 | [6 empty] | 38.9377
| 1 | 5 | 38.07 | 44.1067 | 40.4 | 41.62 | 36.7 | 40 | 40.7 | 39.6 | 41.6 | 37.4 | 38.5 | 39.7 | 39.3279 | 37.88 | 35 | 34.3 | 34.9 | 37.1184 | [6 empty] | 93
| 1 | 6 | 38.25 | 44.037 | 41.2 | 42.21 | 36.4 | 41 | 40.4 | 41.9 | 42.1 | 38.1 | 37.7 | 38.7 | 39.0478 | 38.68 | 36 | 35.5 | 34.8 | 39.2951
| Mean |  | 38.0733 | 43.9262 | 40.6 | 41.5458 | 36 | 40.3333 | 40.95 | 40.45 | 41.0667 | 37.65 | 38.1833 | 40.1 | 39.0875 | 38.6533 | 35.8333 | 35.6833 | 34.7333 | 38.0117
| Median |  | 38.13 | 43.9979 | 41.05 | 41.495 | 36.15 | 40 | 40.65 | 40.3 | 41 | 37.75 | 38.3 | 39.75 | 39.1878 | 38.485 | 36 | 35.55 | 34.75 | 38.206
| Std Dev. |  | 0.22748 | 0.17357 | 0.97775 | 0.37771 | 0.72664 | 0.5164 | 1.1912 | 1.2325 | 0.7174 | 0.85029 | 0.54924 | 1.2696 | 0.95007 | 1.2203 | 0.75277 | 1.0647 | 0.7607 | 1.1927
| Rel.Std.Dev. |  | 0.5975% | 0.3951% | 2.408% | 0.9091% | 2.018% | 1.28% | 2.909% | 3.047% | 1.747% | 2.258% | 1.438% | 3.166% | 2.431% | 3.157% | 2.101% | 2.984% | 2.19% | 3.138%
| PDM3 |  | -2.22% | 12.81% | 4.269% | 6.698% | -7.545% | 3.584% | 5.168% | 3.884% | 5.468% | -3.307% | -1.937% | 2.985% | 0.3847% | -0.7302% | -7.973% | -8.358% | -10.8% | -2.378%
| Z-Score (Absolute) |  | 0.32 | 2.06 | 0.7 | 1.09 | 1.17 | 0.6 | 0.85 | 0.64 | 0.9 | 0.5 | 0.28 | 0.5 | 0.09 | 0.09 | 1.24 | 1.3 | 1.69 | 0.35
| Analytical results for Lu in OREAS 238b (Certified Value 0.11 ppm) | [25 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.21; 1.09: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.11578; 1.06: 0.1; 1.07: 0.106; 1.08: 0.11; 1.09: 0.1574 |  |  |  | 1
| 1 | 2 | 1.01: 0.11713; 1.06: < 0.1; 1.07: < 0.1; 1.08: 0.1; 1.09: 0.1586 |  |  |  | 29
| 1 | 3 | 1.01: 0.12162; 1.06: 0.1; 1.07: 0.103; 1.08: 0.11; 1.09: 0.1601 |  |  |  | 16
| 1 | 4 | 1.01: 0.11953; 1.06: < 0.1; 1.07: 0.101; 1.08: 0.1; 1.09: 0.1453 |  |  |  | 0.10947
| 1 | 5 | 1.01: 0.11822; 1.06: 0.1; 1.07: < 0.1; 1.08: 0.1; 1.09: 0.1738 |  |  |  | 94
| 1 | 6 | 1.01: 0.10641; 1.06: 0.1; 1.07: 0.11; 1.08: 0.1; 1.09: 0.1382
| Mean |  | 1.01: 0.11645; 1.06: 0.1; 1.07: 0.105; 1.08: 0.10333; 1.09: 0.15557
| Median |  | 1.01: 0.11767; 1.06: 0.1; 1.07: 0.1045; 1.08: 0.1; 1.09: 0.158
| Std Dev. |  | 1.01: 0.0053105; 1.06: 0; 1.07: 0.0039158; 1.08: 0.005164; 1.09: 0.01243
| Rel.Std.Dev. |  | 1.01: 4.56%; 1.06: 0%; 1.07: 3.729%; 1.08: 4.997%; 1.09: 7.99%
| PDM3 |  | 1.01: 6.373%; 1.06: -8.652%; 1.07: -4.084%; 1.08: -5.607%; 1.09: 42.11%
| Z-Score (Absolute) |  | 1.01: 0.53; 1.06: 0.81; 1.07: 0.67; 1.08: 0; 1.09: 2.11
| Analytical results for Mg in OREAS 238b (Certified Value 1.41 wt.%) | [25 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 1.267 | 1.33 | 1.45 | 1.4333 | 1.39 | 1.43 | 1.47 | 1.33 | 1.38 | 1.4 | 1.43 | 1.36 | 1.44 | 1.5131 | 1.514 | 1.26 | 1.3526 | 1.46 | 1.45 | 1.5665 | 1.4993 |  |  |  | 1
| 1 | 2 | 1.27 | 1.3401 | 1.43 | 1.4267 | 1.38 | 1.43 | 1.5 | 1.33 | 1.39 | 1.42 | 1.41 | 1.37 | 1.45 | 1.5053 | 1.478 | 1.22 | 1.3627 | 1.44 | 1.41 | 1.535 | 1.5235 |  |  |  | 0
| 1 | 3 | 1.276 | 1.3769 | 1.47 | 1.42 | 1.38 | 1.47 | 1.5 | 1.32 | 1.36 | 1.43 | 1.43 | 1.38 | 1.45 | 1.5154 | 1.508 | 1.27 | 1.3559 | 1.45 | 1.42 | 1.4948 | 1.5443 |  |  |  | 16
| 1 | 4 | 1.28 | 1.3459 | 1.38 | 1.4433 | 1.38 | 1.44 | 1.5 | 1.34 | 1.37 | 1.44 | 1.43 | 1.38 | 1.49 | 1.5268 | 1.441 | 1.26 | 1.3441 | 1.43 | 1.39 | 1.5342 | 1.4536 |  |  |  | 1.4147
| 1 | 5 | 1.263 | 1.3351 | 1.46 | 1.41 | 1.39 | 1.5 | 1.49 | 1.32 | 1.39 | 1.45 | 1.43 | 1.37 | 1.43 | 1.5084 | 1.471 | 1.21 | 1.3651 | 1.44 | 1.38 | 1.5412 | 1.4791 |  |  |  | 95
| 1 | 6 | 1.252 | 1.3546 | 1.47 | 1.4367 | 1.38 | 1.45 | 1.53 | 1.35 | 1.39 | 1.44 | 1.42 | 1.38 | 1.42 | 1.5213 | 1.49 | 1.25 | 1.3819 | 1.45 | 1.4 | 1.533 | 1.4769
| Mean |  | 1.268 | 1.3471 | 1.4433 | 1.4283 | 1.3833 | 1.4533 | 1.4983 | 1.3317 | 1.38 | 1.43 | 1.425 | 1.3733 | 1.4467 | 1.515 | 1.4837 | 1.245 | 1.3604 | 1.445 | 1.4083 | 1.5341 | 1.4961
| Median |  | 1.2685 | 1.343 | 1.455 | 1.43 | 1.38 | 1.445 | 1.5 | 1.33 | 1.385 | 1.435 | 1.43 | 1.375 | 1.445 | 1.5142 | 1.484 | 1.255 | 1.3593 | 1.445 | 1.405 | 1.5346 | 1.4892
| Std Dev. |  | 0.0099398 | 0.016902 | 0.034448 | 0.012065 | 0.005164 | 0.027325 | 0.019408 | 0.01169 | 0.012649 | 0.017889 | 0.0083666 | 0.008165 | 0.024221 | 0.0079959 | 0.026703 | 0.02429 | 0.012943 | 0.010488 | 0.024833 | 0.023007 | 0.033311
| Rel.Std.Dev. |  | 0.7839% | 1.255% | 2.387% | 0.8447% | 0.3733% | 1.88% | 1.295% | 0.8779% | 0.9166% | 1.251% | 0.5871% | 0.5945% | 1.674% | 0.5278% | 1.8% | 1.951% | 0.9515% | 0.7258% | 1.763% | 1.5% | 2.227%
| PDM3 |  | -10.37% | -4.778% | 2.024% | 0.964% | -2.217% | 2.731% | 5.912% | -5.869% | -2.452% | 1.082% | 0.7284% | -2.924% | 2.26% | 7.093% | 4.875% | -12% | -3.841% | 2.142% | -0.4497% | 8.442% | 5.755%
| Z-Score (Absolute) |  | 1.97 | 1 | 0.18 | 0 | 0.55 | 0.31 | 0.86 | 1.19 | 0.59 | 0.02 | 0.04 | 0.67 | 0.22 | 1.06 | 0.68 | 2.25 | 0.83 | 0.2 | 0.25 | 1.3 | 0.83
| Analytical results for Mn in OREAS 238b (Certified Value 0.027 wt.%) | [25 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.02767 | 0.027001 | 0.0276 | 0.029119 | 0.0281 | 0.0263 | 0.0261 | 0.0303 | 0.0267 | 0.0272 | 0.0274 | 0.0254 | 0.0267 | 0.028557 | 0.0256 | 0.0266 | 0.025927 | 0.03084 | 0.0276 | 0.02601 | 0.0317 |  |  |  | 1
| 1 | 2 | 0.02818 | 0.027377 | 0.0269 | 0.029206 | 0.0279 | 0.0263 | 0.0265 | 0.0305 | 0.0268 | 0.0274 | 0.0266 | 0.0257 | 0.0269 | 0.02852 | 0.0256 | 0.0255 | 0.026079 | 0.03056 | 0.0268 | 0.02519 | 0.0322 |  |  |  | 30
| 1 | 3 | 0.02745 | 0.027283 | 0.0266 | 0.029081 | 0.0287 | 0.0267 | 0.0265 | 0.0308 | 0.0265 | 0.0281 | 0.0272 | 0.026 | 0.027 | 0.028599 | 0.0256 | 0.0266 | 0.025975 | 0.03015 | 0.0271 | 0.02469 | 0.0326 |  |  |  | 16
| 1 | 4 | 0.02809 | 0.027406 | 0.0261 | 0.02918 | 0.0281 | 0.0261 | 0.0268 | 0.0307 | 0.0267 | 0.0278 | 0.0271 | 0.0257 | 0.0274 | 0.028675 | 0.024 | 0.0263 | 0.025779 | 0.03099 | 0.0265 | 0.02524 | 0.0311 |  |  |  | 0.027295
| 1 | 5 | 0.0277 | 0.027134 | 0.0285 | 0.028898 | 0.0288 | 0.0266 | 0.0265 | 0.0302 | 0.0269 | 0.0281 | 0.0274 | 0.0255 | 0.0265 | 0.028628 | 0.0248 | 0.0253 | 0.026196 | 0.03114 | 0.0261 | 0.02541 | 0.0316 |  |  |  | 96
| 1 | 6 | 0.02698 | 0.027595 | 0.0278 | 0.029309 | 0.0288 | 0.0266 | 0.0269 | 0.032 | 0.0269 | 0.0276 | 0.0275 | 0.0254 | 0.0261 | 0.0287 | 0.0248 | 0.0263 | 0.026441 | 0.03058 | 0.0266 | 0.0253 | 0.0315
| Mean |  | 0.027678 | 0.027299 | 0.02725 | 0.029132 | 0.0284 | 0.026433 | 0.02655 | 0.03075 | 0.02675 | 0.0277 | 0.0272 | 0.025617 | 0.026767 | 0.028613 | 0.025067 | 0.0261 | 0.026066 | 0.03071 | 0.026783 | 0.025307 | 0.031783
| Median |  | 0.027685 | 0.02733 | 0.02725 | 0.029149 | 0.0284 | 0.02645 | 0.0265 | 0.0306 | 0.02675 | 0.0277 | 0.0273 | 0.0256 | 0.0268 | 0.028613 | 0.0252 | 0.0263 | 0.026027 | 0.03071 | 0.0267 | 0.02527 | 0.03165
| Std Dev. |  | 0.00043852 | 0.00021035 | 0.00087807 | 0.00013898 | 0.00040988 | 0.00023381 | 0.00028107 | 0.00065345 | 0.00015166 | 0.00036878 | 0.00032863 | 0.00023166 | 0.00044572 | 6.8811e-05 | 0.0006532 | 0.00056214 | 0.00023137 | 0.00035587 | 0.00051929 | 0.00042514 | 0.00053448
| Rel.Std.Dev. |  | 1.584% | 0.7705% | 3.222% | 0.4771% | 1.443% | 0.8845% | 1.059% | 2.125% | 0.5669% | 1.331% | 1.208% | 0.9043% | 1.665% | 0.2405% | 2.606% | 2.154% | 0.8876% | 1.159% | 1.939% | 1.68% | 1.682%
| PDM3 |  | 1.405% | 0.01626% | -0.1642% | 6.731% | 4.049% | -3.156% | -2.729% | 12.66% | -1.996% | 1.484% | -0.3474% | -6.148% | -1.935% | 4.83% | -8.163% | -4.377% | -4.501% | 12.51% | -1.874% | -7.284% | 16.44%
| Z-Score (Absolute) |  | 0.29 | 0.06 | 0.03 | 1.18 | 0.74 | 0.47 | 0.4 | 2.18 | 0.28 | 0.31 | 0 | 0.97 | 0.27 | 0.87 | 1.32 | 0.67 | 0.7 | 2.15 | 0.26 | 1.16 | 2.81
| Analytical results for Mo in OREAS 238b (Certified Value 0.81 ppm) | [25 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.82 | 0.81 | < 2 | 0.8 | 0.51 | 0.8 | 0.8 | 0.83 | 0.82 | 0.79 | 0.81 | 0.81 | 0.85189 | < 1 | 1.029 | 0.84 | < 1 | 0.85 | 0.6199 | 2.7469 |  |  |  |  | 1
| 1 | 2 | 0.84 | 0.75 | < 2 | 0.8 | 0.52 | 0.8 | 0.8 | 0.81 | 0.79 | 0.72 | 0.83 | 0.77 | 0.84676 | < 1 | 1.142 | 0.69 | < 1 | 0.85 | 0.5823 | 2.9368 |  |  |  |  | 31
| 1 | 3 | 0.84 | 0.75 | < 2 | 0.8 | 0.68 | 0.8 | 0.8 | 0.85 | 0.83 | 0.72 | 0.78 | 0.86 | 0.86797 | < 1 | 1.08 | 0.79 | < 1 | 0.86 | 0.5797 | 2.7203 |  |  |  |  | 16
| 1 | 4 | 0.82 | 0.75 | < 2 | 0.8 | 0.5 | 0.8 | 0.8 | 0.84 | 0.86 | 0.75 | 0.85 | 0.89 | 0.85119 | < 1 | 0.999 | 0.77 | < 1 | 0.81 | 0.7094 | 2.4485 |  |  |  |  | 0.81068
| 1 | 5 | 0.89 | 0.75 | < 2 | 0.8 | 0.57 | 0.9 | 0.8 | 0.82 | 0.8 | 0.75 | 0.79 | 0.86 | 0.81742 | < 1 | 1.02 | 0.86 | < 1 | 0.83 | 0.6393 | 2.3125 |  |  |  |  | 97
| 1 | 6 | 0.84 | 0.74 | < 2 | 0.8 | 0.63 | 0.9 | 0.9 | 0.85 | 0.8 | 0.7 | 0.79 | 0.85 | 0.81959 | < 1 | 1.026 | 0.74 | < 1 | 0.85 | 0.5797 | 2.1766
| Mean |  | 0.84167 | 0.75833 |  | 0.8 | 0.56833 | 0.83333 | 0.81667 | 0.83333 | 0.81667 | 0.73833 | 0.80833 | 0.84 | 0.84247 |  | 1.0493 | 0.78167 |  | 0.84167 | 0.61838 | 2.5569
| Median |  | 0.84 | 0.75 |  | 0.8 | 0.545 | 0.8 | 0.8 | 0.835 | 0.81 | 0.735 | 0.8 | 0.855 | 0.84897 |  | 1.0275 | 0.78 |  | 0.85 | 0.6011 | 2.5844
| Std Dev. |  | 0.025626 | 0.025626 |  | 1.2162e-16 | 0.073052 | 0.05164 | 0.040825 | 0.01633 | 0.02582 | 0.031885 | 0.027142 | 0.042895 | 0.01992 |  | 0.052686 | 0.063061 |  | 0.018348 | 0.051027 | 0.29094
| Rel.Std.Dev. |  | 3.045% | 3.379% |  | 1.52e-14% | 12.85% | 6.197% | 4.999% | 1.96% | 3.162% | 4.319% | 3.358% | 5.107% | 2.365% |  | 5.021% | 8.067% |  | 2.18% | 8.252% | 11.38%
| PDM3 |  | 3.823% | -6.457% |  | -1.317% | -29.89% | 2.795% | 0.7388% | 2.795% | 0.7388% | -8.924% | -0.2891% | 3.617% | 3.922% |  | 29.44% | -3.579% |  | 3.823% | -23.72% | 215.4%
| Z-Score (Absolute) |  | 0.55 | 1.29 | 4.07 | 0.37 | 5.51 | 0.37 | 0 | 0.37 | 0 | 1.74 | 0.18 | 0.52 | 0.57 | 7.02 | 5.16 | 0.78 | 7.02 | 0.55 | 4.4 | 38.6
| Analytical results for Na in OREAS 238b (Certified Value 0.082 wt.%) | [25 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.083 | 0.11955 | 0.1 | 0.08 | 0.09 | 0.09 | 0.1 | 0.11 | 0.08 | 0.08 | 0.09 | 0.07 | 0.084 | 0.079442 | 0.074 | 0.08 | 0.085024 | 0.07 | 0.09 | 0.0735 | 0.0606 |  |  |  | 1
| 1 | 2 | 0.083 | 0.1213 | 0.1 | 0.08 | 0.09 | 0.09 | 0.11 | 0.12 | 0.08 | 0.08 | 0.08 | 0.07 | 0.084 | 0.082314 | 0.074 | 0.07 | 0.085056 | 0.07 | 0.09 | 0.0713 | 0.06 |  |  |  | 0
| 1 | 3 | 0.084 | 0.12461 | 0.096 | 0.08 | 0.09 | 0.09 | 0.1 | 0.11 | 0.08 | 0.09 | 0.08 | 0.07 | 0.085 | 0.078675 | 0.074 | 0.08 | 0.083134 | 0.07 | 0.09 | 0.0694 | 0.0617 |  |  |  | 16
| 1 | 4 | 0.082 | 0.11997 | 0.094 | 0.08 | 0.09 | 0.09 | 0.11 | 0.11 | 0.08 | 0.09 | 0.09 | 0.08 | 0.087 | 0.083286 | 0.074 | 0.08 | 0.083929 | 0.07 | 0.09 | 0.0718 | 0.0587 |  |  |  | 0.082073
| 1 | 5 | 0.083 | 0.12109 | 0.094 | 0.073333 | 0.09 | 0.09 | 0.11 | 0.13 | 0.08 | 0.09 | 0.09 | 0.07 | 0.083 | 0.079354 | 0.074 | 0.07 | 0.088053 | 0.07 | 0.08 | 0.0719 | 0.059 |  |  |  | 98
| 1 | 6 | 0.081 | 0.12377 | 0.094 | 0.083333 | 0.09 | 0.09 | 0.11 | 0.13 | 0.08 | 0.09 | 0.08 | 0.08 | 0.082 | 0.077209 | 0.074 | 0.08 | 0.088625 | 0.07 | 0.09 | 0.0713 | 0.0594
| Mean |  | 0.082667 | 0.12171 | 0.096333 | 0.079444 | 0.09 | 0.09 | 0.10667 | 0.11833 | 0.08 | 0.086667 | 0.085 | 0.073333 | 0.084167 | 0.080047 | 0.074 | 0.076667 | 0.085637 | 0.07 | 0.088333 | 0.071533 | 0.0599
| Median |  | 0.083 | 0.1212 | 0.095 | 0.08 | 0.09 | 0.09 | 0.11 | 0.115 | 0.08 | 0.09 | 0.085 | 0.07 | 0.084 | 0.079398 | 0.074 | 0.08 | 0.08504 | 0.07 | 0.09 | 0.07155 | 0.0597
| Std Dev. |  | 0.0010328 | 0.0020451 | 0.0029439 | 0.0032773 | 1.5202e-17 | 1.5202e-17 | 0.005164 | 0.0098319 | 0 | 0.005164 | 0.0054772 | 0.005164 | 0.0017224 | 0.0022985 | 0 | 0.005164 | 0.0022208 | 0 | 0.0040825 | 0.0013216 | 0.0011171
| Rel.Std.Dev. |  | 1.249% | 1.68% | 3.056% | 4.125% | 1.689e-14% | 1.689e-14% | 4.841% | 8.309% | 0% | 5.958% | 6.444% | 7.042% | 2.046% | 2.871% | 0% | 6.736% | 2.593% | 0% | 4.622% | 1.848% | 1.865%
| PDM3 |  | 0.7238% | 48.3% | 17.38% | -3.202% | 9.659% | 9.659% | 29.97% | 44.18% | -2.525% | 5.598% | 3.567% | -10.65% | 2.551% | -2.469% | -9.836% | -6.587% | 4.343% | -14.71% | 7.628% | -12.84% | -27.02%
| Z-Score (Absolute) |  | 0.17 | 4.34 | 1.41 | 0.55 | 0.67 | 0.67 | 2.6 | 3.95 | 0.48 | 0.29 | 0.1 | 1.25 | 0 | 0.48 | 1.15 | 0.87 | 0.17 | 1.64 | 0.48 | 1.46 | 2.81
| Analytical results for Nb in OREAS 238b (Certified Value 0.39 ppm) | [25 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | 0.53 | 0.9 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.36 | 0.37 | 0.29 | 0.42 | 0.55993 | 0.51 | 0.28 | 0.3042 | [9 empty] | 1
| 1 | 2 | 0.52 | 0.9 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.38 | 0.34 | 0.28 | 0.38 | 0.59495 | 0.49 | 0.27 | 0.2601 | [9 empty] | 32
| 1 | 3 | 0.49 | 0.8 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.37 | 0.37 | 0.29 | 0.39 | 0.52397 | 0.47 | 0.28 | 0.2418 | [9 empty] | 16
| 1 | 4 | 0.52 | 0.8 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.38 | 0.39 | 0.27 | 0.35 | 0.57203 | 0.5 | 0.26 | 0.2713 | [9 empty] | 0.39277
| 1 | 5 | 0.56 | 0.5 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.37 | 0.34 | 0.3 | 0.38 | 0.58242 | 0.49 | 0.27 | 0.2756 | [9 empty] | 99
| 1 | 6 | 0.49 | 0.5 | < 5 | 0.7 | < 1 | 0.4 | 0.2 | 0.42 | 0.38 | 0.3 | 0.35 | 0.58625 | 0.54 | 0.3 | 0.2452
| Mean |  | 0.51833 | 0.73333 |  | 0.7 |  | 0.4 | 0.2 | 0.38 | 0.365 | 0.28833 | 0.37833 | 0.56992 | 0.5 | 0.27667 | 0.26637
| Median |  | 0.52 | 0.8 |  | 0.7 |  | 0.4 | 0.2 | 0.375 | 0.37 | 0.29 | 0.38 | 0.57723 | 0.495 | 0.275 | 0.2657
| Std Dev. |  | 0.026394 | 0.18619 |  | 1.2162e-16 |  | 6.0809e-17 | 3.0405e-17 | 0.020976 | 0.020736 | 0.01169 | 0.026394 | 0.025552 | 0.023664 | 0.013663 | 0.022942
| Rel.Std.Dev. |  | 5.092% | 25.39% |  | 1.737e-14% |  | 1.52e-14% | 1.52e-14% | 5.52% | 5.681% | 4.054% | 6.977% | 4.483% | 4.733% | 4.938% | 8.613%
| PDM3 |  | 31.97% | 86.71% |  | 78.22% |  | 1.84% | -49.08% | -3.252% | -7.071% | -26.59% | -3.676% | 45.1% | 27.3% | -29.56% | -32.18%
| Z-Score (Absolute) |  | 0.82 | NA | 12.58 | NA | 0.71 | NA | NA | 0 | 0.09 | 0.54 | 0.01 | 1.13 | 0.71 | 0.61 | 0.67
Results from laboratories 1.06, 1.08, 1.10 and 1.11 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 238b (Indicative Value 23.1 ppm) | [25 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 1
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 1
| 1 | 1 | 1.01: 27.808; 1.06: 21.6; 1.07: 25.07; 1.08: 18.6435 |  |  |  | 1
| 1 | 2 | 1.01: 28.3944; 1.06: 21; 1.07: 22.52; 1.08: 17.5349 |  |  |  | 9
| 1 | 3 | 1.01: 27.6084; 1.06: 21.5; 1.07: 25.08; 1.08: 17.2907 |  |  |  | 16
| 1 | 4 | 1.01: 27.3429; 1.06: 20.3; 1.07: 24.59; 1.08: 19.9088 |  |  |  | 23.1412
| 1 | 5 | 1.01: 28.1272; 1.06: 22.3; 1.07: 23.31; 1.08: 20.3325 |  |  |  | 15
| 1 | 6 | 1.01: 28.148; 1.06: 22.6; 1.07: 26.62; 1.08: 17.7602
| Mean |  | 1.01: 27.9048; 1.06: 21.55; 1.07: 24.5317; 1.08: 18.5784
| Median |  | 1.01: 27.9676; 1.06: 21.55; 1.07: 24.83; 1.08: 18.2019
| Std Dev. |  | 1.01: 0.39022; 1.06: 0.84083; 1.07: 1.4487; 1.08: 1.2859
| Rel.Std.Dev. |  | 1.01: 1.398%; 1.06: 3.902%; 1.07: 5.906%; 1.08: 6.921%
| PDM3 |  | 1.01: 20.58%; 1.06: -6.876%; 1.07: 6.009%; 1.08: -19.72%
| Z-Score (Absolute) |  | 1.01: 1.1; 1.06: 0.34; 1.07: 0.34; 1.08: 1.01
| Analytical results for Ni in OREAS 238b (Certified Value 61 ppm) | [25 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  | 1
| 1 | 1 | 58.2 | 53.01 | 60 | 55.87 | 61 | 63 | 64 | 64.8 | 60.9 | 59.8 | 64.2 | 59.2 | 61.3 | 64.8038 | 61 | 44.05 | 62.2 | 74.1 | 57.29 | 65.5 | 56.05 | 58.692 | 61.8728 |  | 1
| 1 | 2 | 58.6 | 53.8365 | 58.7 | 56.2833 | 60 | 65 | 65 | 63.7 | 61.2 | 61.8 | 61.9 | 60 | 61 | 64.2986 | 60 | 43.53 | 56.4 | 67.71 | 57.43 | 64.3 | 54.91 | 58.6773 | 61.066 |  | 34
| 1 | 3 | 59.1 | 53.9835 | 60.9 | 56.03 | 59 | 65 | 66 | 63.8 | 60.2 | 61.6 | 62.8 | 57.9 | 62.4 | 65.4856 | 62 | 46.01 | 61 | 68.06 | 57.41 | 65 | 54.055 | 59.2587 | 62.2518 |  | 16
| 1 | 4 | 56.9 | 53.789 | 57.5 | 56.4267 | 60 | 63 | 68 | 65.9 | 59.1 | 62.8 | 61 | 60.6 | 62.8 | 62.7072 | 59 | 45.38 | 59.2 | 67.57 | 56.53 | 62.9 | 54.34 | 53.9068 | 61.2832 |  | 60.6564
| 1 | 5 | 58.1 | 54.112 | 60.5 | 54.49 | 60 | 64 | 65 | 64.7 | 61.4 | 59.9 | 62.2 | 58.3 | 60.7 | 63.3031 | 61 | 43.96 | 59.5 | 71.04 | 57.38 | 62.7 | 54.91 | 54.2686 | 61.472 |  | 100
| 1 | 6 | 56.5 | 54.302 | 59.6 | 56.81 | 61 | 64 | 66 | 65.5 | 62 | 61.5 | 61.8 | 57.8 | 59.8 | 63.0103 | 61 | 46.23 | 57.6 | 71.94 | 57.25 | 64.4 | 54.815 | 55.7807 | 63.0037
| Mean |  | 57.9 | 53.8388 | 59.5333 | 55.985 | 60.1667 | 64 | 65.6667 | 64.7333 | 60.8 | 61.2333 | 62.3167 | 58.9667 | 61.3333 | 63.9348 | 60.6667 | 44.86 | 59.3167 | 70.07 | 57.215 | 64.1333 | 54.8467 | 56.764 | 61.8249
| Median |  | 58.15 | 53.91 | 59.8 | 56.1567 | 60 | 64 | 65.5 | 64.75 | 61.05 | 61.55 | 62.05 | 58.75 | 61.15 | 63.8009 | 61 | 44.715 | 59.35 | 69.55 | 57.335 | 64.35 | 54.8625 | 57.229 | 61.6724
| Std Dev. |  | 1.002 | 0.44711 | 1.2533 | 0.80187 | 0.75277 | 0.89443 | 1.3663 | 0.88242 | 1.0218 | 1.1673 | 1.0926 | 1.1605 | 1.1094 | 1.1003 | 1.0328 | 1.158 | 2.1283 | 2.7033 | 0.3428 | 1.1219 | 0.68418 | 2.4066 | 0.71668
| Rel.Std.Dev. |  | 1.731% | 0.8305% | 2.105% | 1.432% | 1.251% | 1.398% | 2.081% | 1.363% | 1.681% | 1.906% | 1.753% | 1.968% | 1.809% | 1.721% | 1.702% | 2.581% | 3.588% | 3.858% | 0.5991% | 1.749% | 1.247% | 4.24% | 1.159%
| PDM3 |  | -4.544% | -11.24% | -1.851% | -7.701% | -0.8073% | 5.512% | 8.26% | 6.721% | 0.2368% | 0.9512% | 2.737% | -2.786% | 1.116% | 5.405% | 0.017% | -26.04% | -2.209% | 15.52% | -5.674% | 5.732% | -9.578% | -6.417% | 1.927%
| Z-Score (Absolute) |  | 0.57 | 1.41 | 0.23 | 0.97 | 0.1 | 0.69 | 1.03 | 0.84 | 0.03 | 0.12 | 0.34 | 0.35 | 0.14 | 0.67 | 0 | 3.26 | 0.28 | 1.94 | 0.71 | 0.72 | 1.2 | 0.81 | 0.24
| Analytical results for P in OREAS 238b (Certified Value 0.058 wt.%) | [25 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.0547 | 0.054 | 0.058481 | 0.057 | 0.0578 | 0.0565 | 0.061 | 0.058 | 0.058 | 0.059 | 0.058 | 0.058 | 0.059232 | 0.055 | 0.05 | 0.058255 | 0.06 | 0.0584 | 0.0648 | 0.0645 |  |  |  |  | 1
| 1 | 2 | 0.0552 | 0.054 | 0.058736 | 0.058 | 0.0573 | 0.0576 | 0.061 | 0.058 | 0.058 | 0.058 | 0.059 | 0.058 | 0.059184 | 0.0545 | 0.05 | 0.054051 | 0.06 | 0.0574 | 0.064 | 0.0644 |  |  |  |  | 0
| 1 | 3 | 0.0548 | 0.054 | 0.058915 | 0.058 | 0.0592 | 0.0573 | 0.059 | 0.057 | 0.06 | 0.059 | 0.059 | 0.059 | 0.05898 | 0.0589 | 0.05 | 0.057159 | 0.06 | 0.0574 | 0.0629 | 0.0654 |  |  |  |  | 16
| 1 | 4 | 0.0545 | 0.052 | 0.05852 | 0.058 | 0.0581 | 0.0578 | 0.062 | 0.058 | 0.059 | 0.059 | 0.059 | 0.061 | 0.056636 | 0.0537 | 0.05 | 0.059077 | 0.06 | 0.0563 | 0.0639 | 0.0622 |  |  |  |  | 0.057782
| 1 | 5 | 0.0544 | 0.056 | 0.058954 | 0.057 | 0.0591 | 0.0576 | 0.06 | 0.058 | 0.059 | 0.059 | 0.059 | 0.058 | 0.060181 | 0.055 | 0.05 | 0.057289 | 0.06 | 0.0559 | 0.0637 | 0.0629 |  |  |  |  | 101
| 1 | 6 | 0.0548 | 0.056 | 0.058602 | 0.058 | 0.06 | 0.0581 | 0.062 | 0.058 | 0.058 | 0.058 | 0.059 | 0.057 | 0.05949 | 0.0541 | 0.05 | 0.056824 | 0.06 | 0.0565 | 0.0641 | 0.0628
| Mean |  | 0.054733 | 0.054333 | 0.058701 | 0.057667 | 0.058583 | 0.057483 | 0.060833 | 0.057833 | 0.058667 | 0.058667 | 0.058833 | 0.0585 | 0.05895 | 0.0552 | 0.05 | 0.057109 | 0.06 | 0.056983 | 0.0639 | 0.0637
| Median |  | 0.05475 | 0.054 | 0.058669 | 0.058 | 0.0586 | 0.0576 | 0.061 | 0.058 | 0.0585 | 0.059 | 0.059 | 0.058 | 0.059208 | 0.05475 | 0.05 | 0.057224 | 0.06 | 0.05695 | 0.06395 | 0.06365
| Std Dev. |  | 0.00028048 | 0.0015055 | 0.00020089 | 0.0005164 | 0.0010147 | 0.00054924 | 0.001169 | 0.00040825 | 0.0008165 | 0.0005164 | 0.00040825 | 0.0013784 | 0.001208 | 0.0018826 | 7.6012e-18 | 0.0017119 | 0 | 0.0009196 | 0.00061644 | 0.0012426
| Rel.Std.Dev. |  | 0.5124% | 2.771% | 0.3422% | 0.8955% | 1.732% | 0.9555% | 1.922% | 0.7059% | 1.392% | 0.8802% | 0.6939% | 2.356% | 2.049% | 3.41% | 1.52e-14% | 2.998% | 0% | 1.614% | 0.9647% | 1.951%
| PDM3 |  | -5.277% | -5.969% | 1.59% | -0.2002% | 1.386% | -0.5175% | 5.28% | 0.08821% | 1.53% | 1.53% | 1.819% | 1.242% | 2.021% | -4.469% | -13.47% | -1.165% | 3.838% | -1.383% | 10.59% | 10.24%
| Z-Score (Absolute) |  | 2.06 | 2.28 | 0.09 | 0.47 | 0.02 | 0.57 | 1.24 | 0.38 | 0.07 | 0.07 | 0.16 | 0.02 | 0.22 | 1.81 | 4.62 | 0.78 | 0.79 | 0.84 | 2.9 | 2.79
| Analytical results for Pb in OREAS 238b (Certified Value 14.2 ppm) | [25 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g |  |  | 1
| 1 | 1 | 14.6 | 15.3 | 13.47 | 14.4 | 11.5 | 13 | 16.1 | 14.6 | 14 | 15.4 | 14.8 | 16.01 | 15.0064 | 12 | 26.6 | 14.7 | 6.27 | 12.3 | 13.1 | 16.6 | 16 | 10.4446 |  |  | 1
| 1 | 2 | 14.7 | 15.2 | 13.47 | 14.8 | 11.8 | 12.9 | 15.2 | 14.7 | 14.2 | 15 | 15 | 13.98 | 15.0531 | 13 | 28.4 | 13.5 | 9.43 | 13.7 | 12.3 | 16.3 | 16 | 10.7402 |  |  | 35
| 1 | 3 | 14.4 | 15.6 | 14.2133 | 14.9 | 12.3 | 12.4 | 15.2 | 14.9 | 13.2 | 15.2 | 14.4 | 14.81 | 15.529 | 13 | 25.58 | 14.4 | 6.91 | 13.7 | 12.5 | 15.8 | 17 | 10.3485 |  |  | 16
| 1 | 4 | 14.5 | 14.7 | 13.8033 | 14.5 | 12.3 | 12.4 | 17.6 | 14.3 | 14 | 14.4 | 15.1 | 14.91 | 15.735 | 12 | 22.39 | 13.8 | 6.49 | 12.1 | 12.3 | 16.2 | 17 | 9.1491 |  |  | 14.1612
| 1 | 5 | 14.8 | 15.7 | 12.8167 | 15.4 | 11.3 | 12.4 | 15.5 | 14.7 | 14.1 | 15.3 | 15.2 | 14.65 | 16.0778 | 13 | 36.4 | 14.3 | 2.9 | 14 | 11.9 | 16.5 | 17 | 9.3637 |  |  | 102
| 1 | 6 | 14.3 | 15.7 | 14.0233 | 15.4 | 11.6 | 13 | 18.7 | 14.9 | 13.3 | 15 | 14.6 | 13.9 | 15.4374 | 13 | 26.96 | 14.5 | 10.66 | 12.3 | 12.4 | 16.7 | 18 | 9.1071
| Mean |  | 14.55 | 15.3667 | 13.6328 | 14.9 | 11.8 | 12.6833 | 16.3833 | 14.6833 | 13.8 | 15.05 | 14.85 | 14.71 | 15.4731 | 12.6667 | 27.7217 | 14.2 | 7.11 | 13.0167 | 12.4167 | 16.35 | 16.8333 | 9.8589
| Median |  | 14.55 | 15.45 | 13.6367 | 14.85 | 11.7 | 12.65 | 15.8 | 14.7 | 14 | 15.1 | 14.9 | 14.73 | 15.4832 | 13 | 26.78 | 14.35 | 6.7 | 13 | 12.35 | 16.4 | 17 | 9.8561
| Std Dev. |  | 0.18708 | 0.38816 | 0.49753 | 0.42895 | 0.41952 | 0.31252 | 1.4497 | 0.22286 | 0.43359 | 0.35637 | 0.30822 | 0.76585 | 0.40814 | 0.5164 | 4.7037 | 0.45607 | 2.7144 | 0.86814 | 0.392 | 0.32711 | 0.75277 | 0.73126
| Rel.Std.Dev. |  | 1.286% | 2.526% | 3.65% | 2.879% | 3.555% | 2.464% | 8.849% | 1.518% | 3.142% | 2.368% | 2.076% | 5.206% | 2.638% | 4.077% | 16.97% | 3.212% | 38.18% | 6.669% | 3.157% | 2.001% | 4.472% | 7.417%
| PDM3 |  | 2.745% | 8.512% | -3.732% | 5.217% | -16.67% | -10.44% | 15.69% | 3.687% | -2.551% | 6.276% | 4.864% | 3.875% | 9.264% | -10.55% | 95.76% | 0.2737% | -49.79% | -8.082% | -12.32% | 15.46% | 18.87% | -30.38%
| Z-Score (Absolute) |  | 0.03 | 0.39 | 0.51 | 0.15 | 1.47 | 1.01 | 0.92 | 0.03 | 0.43 | 0.23 | 0.12 | 0.05 | 0.45 | 1.02 | 6.84 | 0.22 | 3.92 | 0.84 | 1.15 | 0.9 | 1.16 | 2.48
| Analytical results for Pd in OREAS 238b (Indicative Value 30.1 ppb) | [25 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.08; 1.07: 1.30 |  |  | ppb
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS |  |  | 1
|  |  | 1.01: 0.5g; 1.06: 01g; 1.07: 0.5g |  |  | 1
| 1 | 1 | 1.01: < 1; 1.06: < 10; 1.07: 35 |  |  | 1
| 1 | 2 | 1.01: < 1; 1.06: < 10; 1.07: 27.8 |  |  | 1
| 1 | 3 | 1.01: < 1; 1.06: < 10; 1.07: 19.7 |  |  | 16
| 1 | 4 | 1.01: < 1; 1.06: < 10; 1.07: 37.7 |  |  | 30.05
| 1 | 5 | 1.01: < 1; 1.06: < 10; 1.07: 34.1 |  |  | 16
| 1 | 6 | 1.01: < 1; 1.06: < 10; 1.07: 26
| Mean |  | 1.07: 30.05
| Median |  | 1.07: 30.95
| Std Dev. |  | 1.07: 6.753
| Rel.Std.Dev. |  | 1.07: 22.47%
PDM3 | bar: 1.07
| Z-Score (Absolute) |  | 1.01: 0.67; 1.06: 0; 1.07: 3.75
| Analytical results for Pr in OREAS 238b (Indicative Value 6.29 ppm) | [25 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  | 2
| 1 | 1 | 1.01: 7.3799; 1.06: 6.3; 1.07: 6.859; 1.08: 4.9621 |  |  | 1
| 1 | 2 | 1.01: 7.4514; 1.06: 6.2; 1.07: 5.963; 1.08: 4.6534 |  |  | 11
| 1 | 3 | 1.01: 7.5455; 1.06: 6.1; 1.07: 6.566; 1.08: 4.6472 |  |  | 16
| 1 | 4 | 1.01: 7.4748; 1.06: 5.8; 1.07: 6.495; 1.08: 5.298 |  |  | 6.2926
| 1 | 5 | 1.01: 7.6354; 1.06: 6.3; 1.07: 6.326; 1.08: 5.431 |  |  | 17
| 1 | 6 | 1.01: 7.535; 1.06: 6.4; 1.07: 6.999; 1.08: 4.7
| Mean |  | 1.01: 7.5037; 1.06: 6.1833; 1.07: 6.5347; 1.08: 4.9486
| Median |  | 1.01: 7.5049; 1.06: 6.25; 1.07: 6.5305; 1.08: 4.8311
| Std Dev. |  | 1.01: 0.088275; 1.06: 0.2137; 1.07: 0.37248; 1.08: 0.34489
| Rel.Std.Dev. |  | 1.01: 1.176%; 1.06: 3.456%; 1.07: 5.7%; 1.08: 6.97%
| PDM3 |  | 1.01: 19.25%; 1.06: -1.736%; 1.07: 3.847%; 1.08: -21.36%
| Z-Score (Absolute) |  | 1.01: 1.17; 1.06: 0.18; 1.07: 0.18; 1.08: 1.44
| Analytical results for Pt in OREAS 238b (Indicative Value 2.37 ppb) | [25 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.08; 1.07: 1.30 |  |  | ppb
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 01g; 1.07: 0.5g |  |  | 2
| 1 | 1 | 1.01: 2; 1.06: < 5; 1.07: 2.3 |  |  | 1
| 1 | 2 | 1.01: 2; 1.06: < 5; 1.07: 3.7 |  |  | 1
| 1 | 3 | 1.01: 2; 1.06: < 5; 1.07: 2.1 |  |  | 16
| 1 | 4 | 1.01: 2; 1.06: < 5; 1.07: 2.4 |  |  | 2.3667
| 1 | 5 | 1.01: 2; 1.06: < 5; 1.07: 4 |  |  | 18
| 1 | 6 | 1.01: 2; 1.06: < 5; 1.07: 1.9
| Mean |  | 1.01: 2; 1.07: 2.7333
| Median |  | 1.01: 2; 1.07: 2.35
| Std Dev. |  | 1.01: 0; 1.07: 0.88694
| Rel.Std.Dev. |  | 1.01: 0%; 1.07: 32.45%
| PDM3 |  | 1.01: -15.49%; 1.07: 15.49%
| Z-Score (Absolute) |  | 1.01: 1.44; 1.06: 0; 1.07: 0.67
| Analytical results for Rb in OREAS 238b (Certified Value 83 ppm) | [25 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.20 | 1.21 | 1.27 | 1.30 | 1.34 | [7 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 84.31 | 72.5619 | 67.8 | 84.59 | 78 | 93.7 | 80.4 | 86.2 | 82.9 | 85.7 | 82.4 | 84.4435 | 69.81 | 82.9 | 86.4 | 80.6827 | 86.8807 | [7 empty] | 1
| 1 | 2 | 85.43 | 73.5844 | 66.2 | 84.91 | 76.2 | 97.3 | 80.2 | 87.2 | 82.1 | 83.6 | 83.9 | 85.7989 | 70.15 | 79.3 | 85.3 | 79.1902 | 84.4662 | [7 empty] | 37
| 1 | 3 | 84.5 | 74.1829 | 67.4 | 85.02 | 80.1 | 91.2 | 79.3 | 84.2 | 85.4 | 83.1 | 81.1 | 81.9343 | 70.7 | 85.9 | 86.5 | 76.6605 | 86.7445 | [7 empty] | 16
| 1 | 4 | 83.14 | 73.4852 | 63.8 | 84.5 | 76.6 | 97.1 | 82.1 | 84.1 | 86.2 | 82.5 | 85.4 | 85.2796 | 69.68 | 82.5 | 82.8 | 88.2788 | 84.6742 | [7 empty] | 83.3596
| 1 | 5 | 85.37 | 73.8386 | 66.4 | 85.03 | 79.9 | 91.6 | 80.4 | 87.2 | 84 | 85.5 | 83.3 | 84.2852 | 69.72 | 79.9 | 82.7 | 88.4726 | 84.3672 | [7 empty] | 103
| 1 | 6 | 82.25 | 74.3136 | 65.9 | 84.81 | 81.5 | 97 | 82.6 | 88 | 83.8 | 83 | 81.4 | 80.634 | 69.95 | 82.3 | 84.7 | 76.7771 | 87.921
| Mean |  | 84.1667 | 73.6611 | 66.25 | 84.81 | 78.7167 | 94.65 | 80.8333 | 86.15 | 84.0667 | 83.9 | 82.9167 | 83.7292 | 70.0017 | 82.1333 | 84.7333 | 81.677 | 85.8423
| Median |  | 84.405 | 73.7115 | 66.3 | 84.86 | 78.95 | 95.35 | 80.4 | 86.7 | 83.9 | 83.35 | 82.85 | 84.3644 | 69.88 | 82.4 | 85 | 79.9364 | 85.7093
| Std Dev. |  | 1.2572 | 0.62843 | 1.4053 | 0.22226 | 2.116 | 2.8515 | 1.2533 | 1.6514 | 1.5253 | 1.3638 | 1.6216 | 2.0154 | 0.38259 | 2.3645 | 1.6789 | 5.4058 | 1.5263
| Rel.Std.Dev. |  | 1.494% | 0.8531% | 2.121% | 0.2621% | 2.688% | 3.013% | 1.55% | 1.917% | 1.814% | 1.626% | 1.956% | 2.407% | 0.5465% | 2.879% | 1.981% | 6.618% | 1.778%
| PDM3 |  | 0.9681% | -11.63% | -20.53% | 1.74% | -5.57% | 13.54% | -3.031% | 3.347% | 0.8482% | 0.6482% | -0.5314% | 0.4434% | -16.02% | -1.471% | 1.648% | -2.019% | 2.978%
| Z-Score (Absolute) |  | 0.14 | 3.31 | 5.74 | 0.36 | 1.65 | 3.59 | 0.95 | 0.8 | 0.11 | 0.06 | 0.27 | 0 | 4.51 | 0.52 | 0.33 | 0.67 | 0.69
| Analytical results for Re in OREAS 238b (Certified Value < 0.001 ppm) | [25 empty] | Certified
Re | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.06: 1.06; 1.07: 1.08; 1.08: 1.09; 1.09: 1.10; 1.10: 1.11; 1.11: 1.12; 1.12: 1.13; 1.13: 1.14; 1.14: 1.15; 1.15: 1.27; 1.16: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.09: AR*OES/MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*MS; 1.15: AR*OES/MS; 1.16: AR*MS |  |  | 3
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 01g; 1.08: 0.2g; 1.09: 01g; 1.10: 0.25g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g; 1.14: 0.5g; 1.15: 0.5g; 1.16: 0.5g |  |  | 3
| 1 | 1 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0011 |  |  | 1
| 1 | 2 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0009 |  |  | 38
| 1 | 3 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0011 |  |  | 16
| 1 | 4 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0005 |  |  | < 0.001
| 1 | 5 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: 0.001; 1.12: < 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0012 |  |  | 104
| 1 | 6 | 1.01: < 0.001; 1.06: < 0.001; 1.07: < 0.05; 1.08: < 0.002; 1.09: < 0.05; 1.10: < 0.005; 1.11: < 0.001; 1.12: 0.001; 1.13: < 0.001; 1.14: < 0.001; 1.15: < 0.001; 1.16: 0.0008
| Mean |  | 1.11: 0.001; 1.12: 0.001; 1.16: 0.00093333
| Median |  | 1.11: 0.001; 1.12: 0.001; 1.16: 0.001
| Std Dev. |  | 1.11: 0; 1.16: 0.0002582
| Rel.Std.Dev. |  | 1.11: 0%; 1.16: 27.66%
PDM3
| Z-Score (Absolute) |  | 1.01: 0.67; 1.06: 0.67; 1.07: 98.45; 1.08: 1.35; 1.09: 98.45; 1.10: 7.42; 1.11: 0.34; 1.12: 0.34; 1.13: 0.67; 1.14: 0.67; 1.15: 0.67; 1.16: 1.08
| Analytical results for Rh in OREAS 238b (Indicative Value 0.001 ppm) | [25 empty] | Indicative
Rh | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 3
|  |  | 1.01: 0.5g |  |  | 3
| 1 | 1 | 1.01: 0.0003 |  |  | 1
| 1 | 2 | 1.01: 0.0016 |  |  | 6
| 1 | 3 | 1.01: 0.0002 |  |  | 16
| 1 | 4 | 1.01: 0.0005 |  |  | 0.00066667
| 1 | 5 | 1.01: 0.0006 |  |  | 19
| 1 | 6 | 1.01: 0.0008
| Mean |  | 1.01: 0.00066667
| Median |  | 1.01: 0.00055
| Std Dev. |  | 1.01: 0.00050465
| Rel.Std.Dev. |  | 1.01: 75.7%
| PDM3 |  | 1.01: -4.441e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ru in OREAS 238b (Indicative Value < 0.0009 ppm) | [25 empty] | Indicative
Ru | Round | Replicate | 1.01: Lab |  |  | 1
| No | No | 1.01: 1.30 |  |  | ppm
|  |  | 1.01: AR*MS |  |  | 3
|  |  | 1.01: 0.5g |  |  | 3
| 1 | 1 | 1.01: < 0.0009 |  |  | 1
| 1 | 2 | 1.01: < 0.0084 |  |  | 14
| 1 | 3 | 1.01: < 0.0009 |  |  | 16
| 1 | 4 | 1.01: < 0.0065 |  |  | < 0.0009
| 1 | 5 | 1.01: < 0.0037 |  |  | 20
| 1 | 6 | 1.01: < 0.0056
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for S in OREAS 238b (Certified Value 0.173 wt.%) | [25 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | wt.%
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.16 | 0.18252 | 0.165 | 0.14667 | 0.18 | 0.17 | 0.17 | 0.19 | 0.19 | 0.18 | 0.19 | 0.17 | 0.18 | 0.145 | 0.16 | 0.005833 | 0.18 | 0.18 | 0.2114 | 0.1795 |  |  |  |  | 1
| 1 | 2 | 0.17 | 0.18505 | 0.164 | 0.14933 | 0.18 | 0.17 | 0.17 | 0.19 | 0.19 | 0.18 | 0.18 | 0.17 | 0.17 | 0.144 | 0.16 | 0.005392 | 0.17 | 0.17 | 0.2069 | 0.1754 |  |  |  |  | 23
| 1 | 3 | 0.17 | 0.19052 | 0.168 | 0.14667 | 0.18 | 0.17 | 0.17 | 0.19 | 0.19 | 0.19 | 0.18 | 0.17 | 0.18 | 0.156 | 0.16 | 0.005294 | 0.17 | 0.18 | 0.2002 | 0.1867 |  |  |  |  | 16
| 1 | 4 | 0.17 | 0.18733 | 0.163 | 0.14933 | 0.18 | 0.17 | 0.18 | 0.19 | 0.19 | 0.19 | 0.18 | 0.17 | 0.19 | 0.14 | 0.16 | 0.005508 | 0.17 | 0.17 | 0.2059 | 0.1665 |  |  |  |  | 0.17281
| 1 | 5 | 0.16 | 0.18461 | 0.166 | 0.13333 | 0.18 | 0.17 | 0.17 | 0.19 | 0.19 | 0.19 | 0.18 | 0.17 | 0.18 | 0.144 | 0.16 | 0.005611 | 0.18 | 0.17 | 0.2047 | 0.1613 |  |  |  |  | 105
| 1 | 6 | 0.16 | 0.18614 | 0.171 | 0.14667 | 0.18 | 0.17 | 0.18 | 0.19 | 0.19 | 0.19 | 0.18 | 0.17 | 0.18 | 0.145 | 0.16 | 0.005638 | 0.18 | 0.17 | 0.2034 | 0.1627
| Mean |  | 0.165 | 0.18603 | 0.16617 | 0.14533 | 0.18 | 0.17 | 0.17333 | 0.19 | 0.19 | 0.18667 | 0.18167 | 0.17 | 0.18 | 0.14567 | 0.16 | 0.005546 | 0.175 | 0.17333 | 0.20542 | 0.17202
| Median |  | 0.165 | 0.1856 | 0.1655 | 0.14667 | 0.18 | 0.17 | 0.17 | 0.19 | 0.19 | 0.19 | 0.18 | 0.17 | 0.18 | 0.1445 | 0.16 | 0.0055595 | 0.175 | 0.17 | 0.2053 | 0.17095
| Std Dev. |  | 0.0054772 | 0.0027249 | 0.0029269 | 0.0060222 | 3.0405e-17 | 0 | 0.005164 | 3.0405e-17 | 3.0405e-17 | 0.005164 | 0.0040825 | 0 | 0.0063246 | 0.0053914 | 0 | 0.00019169 | 0.0054772 | 0.005164 | 0.0037424 | 0.01015
| Rel.Std.Dev. |  | 3.32% | 1.465% | 1.761% | 4.144% | 1.689e-14% | 0% | 2.979% | 1.6e-14% | 1.6e-14% | 2.766% | 2.247% | 0% | 3.514% | 3.701% | 0% | 3.456% | 3.13% | 2.979% | 1.822% | 5.901%
| PDM3 |  | -4.518% | 7.651% | -3.843% | -15.9% | 4.162% | -1.625% | 0.3039% | 9.949% | 9.949% | 8.02% | 5.126% | -1.625% | 4.162% | -15.71% | -7.412% | -96.79% | 1.268% | 0.3039% | 18.87% | -0.458%
| Z-Score (Absolute) |  | 0.67 | 1.03 | 0.58 | 2.27 | 0.54 | 0.27 | 0 | 1.35 | 1.35 | 1.08 | 0.67 | 0.27 | 0.54 | 2.24 | 1.08 | 13.58 | 0.13 | 0 | 2.6 | 0.11
| Analytical results for Sb in OREAS 238b (Certified Value 469 ppm) | [25 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES |  |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  | 0
| 1 | 1 | 538.36 | 413 | 511.095 | 426.92 | 262.18 | 391 | 462.95 | 489 | 526 | 535 | 498 | 577.43 | 534.6201 | 381 | 566.8 | 390 | 584.21 | 561.4 | 425.8 | 367.5 | 161.7382 |  |  |  | 1
| 1 | 2 | 567.48 | 385 | 511.87 | 397.31 | 265.47 | 434 | 454.43 | 493 | 515 | 517 | 499 | 562.72 | 543.3616 | 369 | 574.7 | 376 | 512.39 | 560.7 | 398.01 | 353.7 | 162.1188 |  |  |  | 39
| 1 | 3 | 551.39 | 362 | 508.97 | 397.12 | 286.9 | 406 | 464.69 | 511 | 534 | 518 | 509 | 567.42 | 534.0242 | 395 | 577.9 | 402 | 563.79 | 554.2 | 410.29 | 341.6 | 161.7382 |  |  |  | 16
| 1 | 4 | 546.61 | 370 | 513.945 | 415.69 | 278.62 | 406 | 483.63 | 501 | 518 | 512 | 498 | 575.38 | 525.9028 | 361 | 551 | 395 | 569.01 | 554.8 | 406.01 | 352 | 161.7809 |  |  |  | 468.6505
| 1 | 5 | 530.26 | 442 | 516.685 | 416.9 | 282.41 | 434 | 460.57 | 485 | 526 | 533 | 489 | 558.21 | 537.1554 | 375 | 508 | 397 | 546.11 | 563.6 | 406.6 | 346.1 | 161.8547 |  |  |  | 106
| 1 | 6 | 521.36 | 412 | 519.32 | 411.61 | 271.65 | 444 | 468.79 | 501 | 521 | 518 | 489 | 557.07 | 537.7833 | 380 | 568.2 | 389 | 549.93 | 555.1 | 424.47 | 345.5 | 161.9285
| Mean |  | 542.5767 | 397.3333 | 513.6475 | 410.925 | 274.5383 | 419.1667 | 465.8433 | 496.6667 | 523.3333 | 522.1667 | 497 | 566.3717 | 535.4746 | 376.8333 | 557.7667 | 391.5 | 554.24 | 558.3 | 411.8633 | 351.0667 | 161.8599
| Median |  | 542.485 | 398.5 | 512.9075 | 413.65 | 275.135 | 420 | 463.82 | 497 | 523.5 | 518 | 498 | 565.07 | 535.8877 | 377.5 | 567.5 | 392.5 | 556.86 | 557.9 | 408.445 | 349.05 | 161.8178
| Std Dev. |  | 16.3325 | 30.3425 | 3.8222 | 11.753 | 9.7441 | 20.9611 | 9.928 | 9.5009 | 6.802 | 9.4534 | 7.4565 | 8.614 | 5.7413 | 11.6003 | 26.0962 | 8.961 | 24.6725 | 4.0684 | 11.041 | 9.1993 | 0.14671
| Rel.Std.Dev. |  | 3.01% | 7.637% | 0.7441% | 2.86% | 3.549% | 5.001% | 2.131% | 1.913% | 1.3% | 1.81% | 1.5% | 1.521% | 1.072% | 3.078% | 4.679% | 2.289% | 4.452% | 0.7287% | 2.681% | 2.62% | 0.09064%
| PDM3 |  | 15.77% | -15.22% | 9.601% | -12.32% | -41.42% | -10.56% | -0.599% | 5.978% | 11.67% | 11.42% | 6.049% | 20.85% | 14.26% | -19.59% | 19.02% | -16.46% | 18.26% | 19.13% | -12.12% | -25.09% | -65.46%
| Z-Score (Absolute) |  | 0.5 | 1.09 | 0.19 | 0.94 | 2.43 | 0.85 | 0.34 | 0 | 0.29 | 0.28 | 0 | 0.76 | 0.42 | 1.31 | 0.67 | 1.15 | 0.63 | 0.67 | 0.93 | 1.59 | 3.66
| Analytical results for Sc in OREAS 238b (Certified Value 6.24 ppm) | [25 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 6.39 | 5.9031 | 5.9 | 6.07 | 6 | 7 | 7 | 6.2 | 7.4 | 6.8 | 6.2 | 6.9 | 5.8 | 6.6597 | 6 | 6.239 | 6.6 | 6.3 | 6.3 | 6.05 | 10.956 | 6.675 |  |  | 1
| 1 | 2 | 6.45 | 5.8789 | 5.8 | 6.0833 | 6 | 6 | 7 | 6.2 | 7.4 | 7 | 6.2 | 6.9 | 5.9 | 6.6477 | 5 | 5.684 | 6.2 | 6.3 | 6 | 5.88 | 10.9982 | 6.3918 |  |  | 40
| 1 | 3 | 6.44 | 5.8908 | 6 | 6.0567 | 6 | 6 | 7 | 6.2 | 7.2 | 7.2 | 6.2 | 6.6 | 5.9 | 6.8747 | 6 | 5.797 | 7 | 6.2 | 6.1 | 5.77 | 11.2049 | 6.6401 |  |  | 16
| 1 | 4 | 6.37 | 5.7308 | 6 | 6.1367 | 6 | 6 | 7 | 6.3 | 7.3 | 7.4 | 6.1 | 7.2 | 6 | 6.9632 | 6 | 5.79 | 6.4 | 6 | 6 | 5.91 | 10.6276 | 6.4812 |  |  | 6.2439
| 1 | 5 | 6.44 | 5.888 | 6.1 | 5.6233 | 6 | 6 | 7 | 6.2 | 7.5 | 6.9 | 6.3 | 7 | 5.8 | 6.8508 | 5 | 5.587 | 6.4 | 6.3 | 5.8 | 5.93 | 10.7905 | 6.4536 |  |  | 107
| 1 | 6 | 6.46 | 5.7511 | 5.7 | 6.1833 | 6 | 6 | 7 | 6.4 | 7.6 | 7.1 | 6.1 | 6.7 | 5.7 | 6.8068 | 5 | 5.825 | 6.4 | 6.3 | 6 | 5.93 | 10.7066 | 6.5196
| Mean |  | 6.425 | 5.8405 | 5.9167 | 6.0256 | 6 | 6.1667 | 7 | 6.25 | 7.4 | 7.0667 | 6.1833 | 6.8833 | 5.85 | 6.8005 | 5.5 | 5.8203 | 6.5 | 6.2333 | 6.0333 | 5.9117 | 10.8806 | 6.5269
| Median |  | 6.44 | 5.8835 | 5.95 | 6.0767 | 6 | 6 | 7 | 6.2 | 7.4 | 7.05 | 6.2 | 6.9 | 5.85 | 6.8288 | 5.5 | 5.7935 | 6.4 | 6.3 | 6 | 5.92 | 10.8733 | 6.5004
| Std Dev. |  | 0.036194 | 0.077702 | 0.1472 | 0.20263 | 0 | 0.40825 | 0 | 0.083666 | 0.14142 | 0.21602 | 0.075277 | 0.2137 | 0.10488 | 0.12467 | 0.54772 | 0.22352 | 0.27568 | 0.12111 | 0.1633 | 0.090425 | 0.21311 | 0.11002
| Rel.Std.Dev. |  | 0.5633% | 1.33% | 2.488% | 3.363% | 0% | 6.62% | 0% | 1.339% | 1.911% | 3.057% | 1.217% | 3.105% | 1.793% | 1.833% | 9.959% | 3.84% | 4.241% | 1.943% | 2.707% | 1.53% | 1.959% | 1.686%
| PDM3 |  | 2.901% | -6.461% | -5.24% | -3.496% | -3.905% | -1.236% | 12.11% | 0.09845% | 18.52% | 13.18% | -0.9693% | 10.24% | -6.308% | 8.915% | -11.91% | -6.783% | 4.102% | -0.1685% | -3.372% | -5.32% | 74.26% | 4.533%
| Z-Score (Absolute) |  | 0.47 | 0.81 | 0.64 | 0.4 | 0.46 | 0.09 | 1.74 | 0.09 | 2.61 | 1.88 | 0.05 | 1.48 | 0.79 | 1.3 | 1.55 | 0.85 | 0.64 | 0.05 | 0.38 | 0.65 | 10.24 | 0.7
| Analytical results for Se in OREAS 238b (Indicative Value 0.29 ppm) | [25 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.17 | 0.2 | < 5 | < 1 | 0.7 | < 1 | 2.4 | 0.4 | 0.2 | 0.2 | 0.2 | 0.6 | < 1 | 0.19 | < 1 | 0.3 | 0.3063 | 3.1936 | 0.36722 | [5 empty] | 1
| 1 | 2 | 0.21 | 0.4 | < 5 | < 1 | 0.8 | < 1 | 2.2 | 0.2 | 0.3 | 0.2 | 0.3 | 0.6 | < 1 | 0.2 | < 1 | 0.2 | 0.218 | 2.9854 | 0.35322 | [5 empty] | 41
| 1 | 3 | 0.2 | 0.3 | < 5 | < 1 | 0.7 | < 1 | 2 | 0.2 | 0.4 | 0.2 | 0.3 | 0.4 | < 1 | 0.21 | < 1 | 0.3 | 0.1862 | 3.0076 | 0.37188 | [5 empty] | 16
| 1 | 4 | 0.18 | 0.4 | < 5 | < 1 | 0.7 | < 1 | 2.3 | 0.4 | 0.3 | 0.2 | 0.3 | 0.8 | < 1 | 0.22 | < 1 | 0.3 | 0.2331 | 2.634 | 0.33289 | [5 empty] | 0.29015
| 1 | 5 | 0.21 | 0.2 | < 5 | < 1 | 0.9 | < 1 | 2.4 | 0.3 | 0.2 | 0.3 | 0.4 | 0.4 | < 1 | 0.23 | < 1 | 0.3 | 0.2828 | 3.1055 | 0.32064 | [5 empty] | 14
| 1 | 6 | 0.2 | 0.3 | < 5 | < 1 | 0.8 | < 1 | 2.4 | 0.4 | 0.3 | 0.2 | 0.3 | 0.5 | < 1 | 0.24 | < 1 | 0.3 | 0.1821 | 2.9514 | 0.33546
| Mean |  | 0.195 | 0.3 |  |  | 0.76667 |  | 2.2833 | 0.31667 | 0.28333 | 0.21667 | 0.3 | 0.55 |  | 0.215 |  | 0.28333 | 0.23475 | 2.9796 | 0.34688
| Median |  | 0.2 | 0.3 |  |  | 0.75 |  | 2.35 | 0.35 | 0.3 | 0.2 | 0.3 | 0.55 |  | 0.215 |  | 0.3 | 0.22555 | 2.9965 | 0.34434
| Std Dev. |  | 0.016432 | 0.089443 |  |  | 0.08165 |  | 0.16021 | 0.098319 | 0.075277 | 0.040825 | 0.063246 | 0.15166 |  | 0.018708 |  | 0.040825 | 0.050677 | 0.19115 | 0.020463
| Rel.Std.Dev. |  | 8.427% | 29.81% |  |  | 10.65% |  | 7.016% | 31.05% | 26.57% | 18.84% | 21.08% | 27.57% |  | 8.702% |  | 14.41% | 21.59% | 6.415% | 5.899%
| PDM3 |  | -32.79% | 3.395% |  |  | 164.2% |  | 687% | 9.139% | -2.349% | -25.33% | 3.395% | 89.56% |  | -25.9% |  | -2.349% | -19.09% | 926.9% | 19.55%
| Z-Score (Absolute) |  | 0.67 | 0.21 | 9.56 | 0.68 | 1.86 | 0.68 | 8.6 | 0.13 | 0.28 | 0.58 | 0.21 | 0.9 | 0.68 | 0.59 | 0.68 | 0.28 | 0.5 | 11.69 | 0
| Analytical results for Si in OREAS 238b (Indicative Value 0.395 wt.%) | [25 empty] | Indicative
Si | Round | Replicate | 1.01: Lab |  |  |  | 1
| No | No | 1.01: 1.30 |  |  |  | wt.%
|  |  | 1.01: AR*OES |  |  |  | 3
|  |  | 1.01: 0.5g |  |  |  | 3
| 1 | 1 | 1.01: 0.4195 |  |  |  | 1
| 1 | 2 | 1.01: 0.3723 |  |  |  | 9
| 1 | 3 | 1.01: 0.3669 |  |  |  | 16
| 1 | 4 | 1.01: 0.4015 |  |  |  | 0.39457
| 1 | 5 | 1.01: 0.4012 |  |  |  | 15
| 1 | 6 | 1.01: 0.406
| Mean |  | 1.01: 0.39457
| Median |  | 1.01: 0.40135
| Std Dev. |  | 1.01: 0.020522
| Rel.Std.Dev. |  | 1.01: 5.201%
| PDM3 |  | 1.01: -8.882e-14%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sm in OREAS 238b (Indicative Value 4.28 ppm) | [25 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 5.1495; 1.06: 3.1; 1.07: 4.635; 1.08: 3.5279 |  |  |  | 1
| 1 | 2 | 1.01: 5.2714; 1.06: 4.5; 1.07: 4.159; 1.08: 3.4162 |  |  |  | 10
| 1 | 3 | 1.01: 5.3864; 1.06: 3.9; 1.07: 4.599; 1.08: 3.3396 |  |  |  | 16
| 1 | 4 | 1.01: 4.992; 1.06: 3.5; 1.07: 4.497; 1.08: 3.9438 |  |  |  | 4.275
| 1 | 5 | 1.01: 5.5698; 1.06: 3.3; 1.07: 4.3; 1.08: 3.8525 |  |  |  | 16
| 1 | 6 | 1.01: 5.3195; 1.06: 3.8; 1.07: 4.905; 1.08: 3.6364
| Mean |  | 1.01: 5.2814; 1.06: 3.6833; 1.07: 4.5158; 1.08: 3.6194
| Median |  | 1.01: 5.2955; 1.06: 3.65; 1.07: 4.548; 1.08: 3.5821
| Std Dev. |  | 1.01: 0.19845; 1.06: 0.49967; 1.07: 0.26338; 1.08: 0.23992
| Rel.Std.Dev. |  | 1.01: 3.757%; 1.06: 13.57%; 1.07: 5.832%; 1.08: 6.629%
| PDM3 |  | 1.01: 23.54%; 1.06: -13.84%; 1.07: 5.633%; 1.08: -15.34%
| Z-Score (Absolute) |  | 1.01: 1.78; 1.06: 0.63; 1.07: 0.63; 1.08: 0.72
| Analytical results for Sn in OREAS 238b (Certified Value 1.45 ppm) | [25 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.5 | 2.6473 | 1.4 | < 10 | 1.4 | 1.4 | 1.6 | 1.4 | 1.4 | 1.4 | 1.4 | 1.5 | 1.544 | < 1 | 1.4 | < 5 | 1.4 | 1.1595 | 2.3089 | [5 empty] | 1
| 1 | 2 | 1.53 | 2.4983 | 1.35 | < 10 | 1.4 | 1.4 | 1.7 | 1.4 | 1.5 | 1.4 | 1.4 | 1.5 | 1.5415 | < 1 | 1.2 | < 5 | 1.4 | 1.1176 | 2.348 | [5 empty] | 26
| 1 | 3 | 1.49 | 2.4818 | 1.32 | < 10 | 1.3 | 1.5 | 1.6 | 1.4 | 1.4 | 1.4 | 1.4 | 1.5 | 1.5301 | < 1 | 1.5 | < 5 | 1.4 | 1.0519 | 2.3247 | [5 empty] | 16
| 1 | 4 | 1.5 | 2.5654 | 1.23 | < 10 | 1.4 | 1.4 | 1.8 | 1.5 | 1.4 | 1.4 | 1.4 | 1.5 | 1.5446 | < 1 | 1.4 | < 5 | 1.3 | 1.2212 | 2.3948 | [5 empty] | 1.446
| 1 | 5 | 1.48 | 2.7152 | 1.44 | < 10 | 1.3 | 1.5 | 1.7 | 1.4 | 1.5 | 1.4 | 1.4 | 1.5 | 1.5224 | < 1 | 1.4 | < 5 | 1.3 | 1.2763 | 2.3435 | [5 empty] | 108
| 1 | 6 | 1.48 | 2.6728 | 1.47 | < 10 | 1.4 | 1.5 | 1.7 | 1.5 | 1.5 | 1.4 | 1.4 | 1.5 | 1.5534 | < 1 | 1.4 | < 5 | 1.3 | 1.0902 | 2.3974
| Mean |  | 1.4967 | 2.5968 | 1.3683 |  | 1.3667 | 1.45 | 1.6833 | 1.4333 | 1.45 | 1.4 | 1.4 | 1.5 | 1.5393 |  | 1.3833 |  | 1.35 | 1.1528 | 2.3529
| Median |  | 1.495 | 2.6063 | 1.375 |  | 1.4 | 1.45 | 1.7 | 1.4 | 1.45 | 1.4 | 1.4 | 1.5 | 1.5427 |  | 1.4 |  | 1.35 | 1.1385 | 2.3457
| Std Dev. |  | 0.018619 | 0.096192 | 0.087502 |  | 0.05164 | 0.054772 | 0.075277 | 0.05164 | 0.054772 | 2.4324e-16 | 2.4324e-16 | 0 | 0.011149 |  | 0.098319 |  | 0.054772 | 0.084041 | 0.036275
| Rel.Std.Dev. |  | 1.244% | 3.704% | 6.395% |  | 3.779% | 3.777% | 4.472% | 3.603% | 3.777% | 1.737e-14% | 1.737e-14% | 0% | 0.7243% |  | 7.107% |  | 4.057% | 7.29% | 1.542%
| PDM3 |  | 3.506% | 79.59% | -5.369% |  | -5.485% | 0.2785% | 16.42% | -0.8742% | 0.2785% | -3.179% | -3.179% | 3.736% | 6.456% |  | -4.332% |  | -6.637% | -20.28% | 62.72%
| Z-Score (Absolute) |  | 0.38 | 9.28 | 0.66 | 28.73 | 0.67 | 0 | 1.89 | 0.13 | 0 | 0.4 | 0.4 | 0.4 | 0.72 | 7.69 | 0.54 | 8.5 | 0.81 | 2.4 | 7.31
| Analytical results for Sr in OREAS 238b (Certified Value 26.1 ppm) | [25 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 26.56 | 27.8231 | 32.9 | 37.6167 | 26.6 | 27 | 31.8 | 26 | 26.8 | 24.7 | 26.4 | 27.4 | 25.9 | 27.2415 | 26 | 23.09 | 24.7 | 24.3 | 27.9 | 20.4 | 120.9362 | 27.5979 |  |  | 1
| 1 | 2 | 26.77 | 26.9847 | 33 | 37.7 | 26.5 | 25 | 32.5 | 26 | 27.1 | 25 | 25.7 | 27.6 | 26.4 | 26.8755 | 25 | 24.8 | 23.8 | 24.3 | 27.1 | 19.6 | 121.3207 | 26.8434 |  |  | 27
| 1 | 3 | 26.93 | 27.7028 | 32.6 | 37.4 | 27.3 | 26 | 31 | 25 | 26.6 | 26.7 | 25.7 | 26.7 | 26.5 | 27.9231 | 26 | 23.72 | 24.6 | 23.7 | 27.7 | 19 | 121.8687 | 25.9793 |  |  | 16
| 1 | 4 | 26.89 | 27.5429 | 31.7 | 37.5233 | 26.7 | 26 | 31.5 | 26 | 26.6 | 26.5 | 24.6 | 28.2 | 27.1 | 27.4254 | 25 | 23.07 | 24.2 | 23.3 | 26.5 | 19.7 | 120.1616 | 25.2074 |  |  | 26.0732
| 1 | 5 | 26.98 | 29.0947 | 33.1 | 36.04 | 26.7 | 26 | 31.8 | 25 | 27.5 | 25.1 | 26.4 | 27.9 | 25.7 | 27.7147 | 25 | 23.1 | 23.3 | 24.4 | 25.7 | 19.8 | 120.402 | 25.668 |  |  | 109
| 1 | 6 | 26.48 | 28.5501 | 32.2 | 38.1167 | 27.3 | 26 | 33.3 | 26 | 27.7 | 25.6 | 25.5 | 26.7 | 25.1 | 27.9614 | 26 | 24.76 | 24.6 | 23.9 | 26.4 | 19.7 | 120.3836 | 26.2789
| Mean |  | 26.7683 | 27.9497 | 32.5833 | 37.3994 | 26.85 | 26 | 31.9833 | 25.6667 | 27.05 | 25.6 | 25.7167 | 27.4167 | 26.1167 | 27.5236 | 25.5 | 23.7567 | 24.2 | 23.9833 | 26.8833 | 19.7 | 120.8455 | 26.2625
| Median |  | 26.83 | 27.763 | 32.75 | 37.57 | 26.7 | 26 | 31.8 | 26 | 26.95 | 25.35 | 25.7 | 27.5 | 26.15 | 27.57 | 25.5 | 23.41 | 24.4 | 24.1 | 26.8 | 19.7 | 120.6691 | 26.1291
| Std Dev. |  | 0.20605 | 0.75437 | 0.54191 | 0.70928 | 0.35637 | 0.63246 | 0.8085 | 0.5164 | 0.46797 | 0.82946 | 0.66758 | 0.61779 | 0.69976 | 0.42357 | 0.54772 | 0.82991 | 0.55498 | 0.43089 | 0.84004 | 0.44721 | 0.65708 | 0.85646
| Rel.Std.Dev. |  | 0.7698% | 2.699% | 1.663% | 1.896% | 1.327% | 2.433% | 2.528% | 2.012% | 1.73% | 3.24% | 2.596% | 2.253% | 2.679% | 1.539% | 2.148% | 3.493% | 2.293% | 1.797% | 3.125% | 2.27% | 0.5437% | 3.261%
| PDM3 |  | 2.666% | 7.197% | 24.97% | 43.44% | 2.979% | -0.2807% | 22.67% | -1.559% | 3.746% | -1.815% | -1.367% | 5.153% | 0.1668% | 5.563% | -2.198% | -8.885% | -7.184% | -8.015% | 3.107% | -24.44% | 363.5% | 0.726%
| Z-Score (Absolute) |  | 0.18 | 1.01 | 4.25 | 7.63 | 0.23 | 0.36 | 3.83 | 0.6 | 0.37 | 0.64 | 0.56 | 0.63 | 0.28 | 0.71 | 0.71 | 1.93 | 1.62 | 1.78 | 0.26 | 4.78 | 66.13 | 0.18
| Analytical results for Ta in OREAS 238b (Certified Value < 0.01 ppm) | [25 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.21 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [10 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.2427 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.1516 | [10 empty] | 28
| 1 | 3 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.03 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0431 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.1625 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0374 | [10 empty] | 110
| 1 | 6 | < 0.01 | < 0.05 | < 5 | < 0.05 | 0.02 | < 0.05 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.0393
| Mean |  | 1.09: 0.021667; 1.12: 0.01; 1.13: 0.01; 1.19: 0.11277
| Median |  | 1.09: 0.02; 1.12: 0.01; 1.13: 0.01; 1.19: 0.09735
| Std Dev. |  | 1.09: 0.0040825; 1.12: 0; 1.13: 0; 1.19: 0.085785
| Rel.Std.Dev. |  | 1.09: 18.84%; 1.12: 0%; 1.13: 0%; 1.19: 76.07%
PDM3
| Z-Score (Absolute) |  | 0.78 | 0.07 | 105.48 | 0.07 | 0.07 | 0.07 | 1.14 | 0.57 | 0.57 | 0.78 | 0.78 | 0.07 | 0.78 | 3.81
| Analytical results for Tb in OREAS 238b (Indicative Value 0.46 ppm) | [25 empty] | Indicative
Tb | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.21; 1.09: 1.30 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*OES |  |  |  | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.5g |  |  |  | 2
| 1 | 1 | 1.01: 0.53483; 1.06: 0.4; 1.07: 0.477; 1.08: 0.45; 1.09: 1.41 |  |  |  | 1
| 1 | 2 | 1.01: 0.54557; 1.06: 0.4; 1.07: 0.432; 1.08: 0.42; 1.09: 1.42 |  |  |  | 11
| 1 | 3 | 1.01: 0.5452; 1.06: 0.4; 1.07: 0.465; 1.08: 0.44; 1.09: 1.36 |  |  |  | 16
| 1 | 4 | 1.01: 0.52182; 1.06: 0.4; 1.07: 0.456; 1.08: 0.41; 1.09: 1.44 |  |  |  | 0.45794
| 1 | 5 | 1.01: 0.5522; 1.06: 0.4; 1.07: 0.438; 1.08: 0.41; 1.09: 1.42 |  |  |  | 17
| 1 | 6 | 1.01: 0.55094; 1.06: 0.4; 1.07: 0.488; 1.08: 0.43; 1.09: 1.38
| Mean |  | 1.01: 0.54176; 1.06: 0.4; 1.07: 0.45933; 1.08: 0.42667; 1.09: 1.405
| Median |  | 1.01: 0.54538; 1.06: 0.4; 1.07: 0.4605; 1.08: 0.425; 1.09: 1.415
| Std Dev. |  | 1.01: 0.011533; 1.06: 6.0809e-17; 1.07: 0.021814; 1.08: 0.01633; 1.09: 0.029496
| Rel.Std.Dev. |  | 1.01: 2.129%; 1.06: 1.52e-14%; 1.07: 4.749%; 1.08: 3.827%; 1.09: 2.099%
| PDM3 |  | 1.01: 18.3%; 1.06: -12.65%; 1.07: 0.3051%; 1.08: -6.828%; 1.09: 206.8%
| Z-Score (Absolute) |  | 1.01: 0.94; 1.06: 0.67; 1.07: 0; 1.08: 0.37; 1.09: 10.75
| Analytical results for Te in OREAS 238b (Indicative Value 0.02 ppm) | [25 empty] | Indicative
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.21 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.02 | 0.06 | 5.7273 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.04 | 0.0478 | [9 empty] | 1
| 1 | 2 | < 0.02 | 0.03 | 5.265 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.02 | < 0.05 | 0.04 | 0.0239 | [9 empty] | 5
| 1 | 3 | < 0.02 | 0.03 | < 5 | < 0.05 | 0.06 | < 0.05 | < 0.1 | 0.02 | 0.03 | 0.01 | 0.01 | < 0.02 | < 0.05 | 0.05 | [10 empty] | 16
| 1 | 4 | 0.02 | 0.05 | 5.7623 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | < 0.05 | 0.04 | 0.012 | [9 empty] | 0.020467
| 1 | 5 | < 0.02 | 0.06 | < 5 | < 0.05 | 0.05 | < 0.05 | < 0.1 | 0.02 | 0.03 | 0.02 | 0.02 | < 0.02 | < 0.05 | 0.05 | [10 empty] | 18
| 1 | 6 | 0.03 | 0.05 | 5.1962 | < 0.05 | 0.04 | < 0.05 | < 0.1 | 0.02 | 0.02 | 0.03 | 0.01 | 0.02 | < 0.05 | 0.04 | 0.0239
| Mean |  | 0.023333 | 0.046667 | 5.4877 |  | 0.048333 |  |  | 0.02 | 0.023333 | 0.02 | 0.016667 | 0.02 |  | 0.043333 | 0.0269
| Median |  | 0.02 | 0.05 | 5.4962 |  | 0.05 |  |  | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 |  | 0.04 | 0.0239
| Std Dev. |  | 0.0057735 | 0.013663 | 0.29857 |  | 0.0075277 |  |  | 0 | 0.005164 | 0.0063246 | 0.005164 | 0 |  | 0.005164 | 0.01502
| Rel.Std.Dev. |  | 24.74% | 29.28% | 5.441% |  | 15.57% |  |  | 0% | 22.13% | 31.62% | 30.98% | 0% |  | 11.92% | 55.84%
| PDM3 |  | 14.01% | 128% | 26713% |  | 136.2% |  |  | -2.28% | 14.01% | -2.28% | -18.57% | -2.28% |  | 111.7% | 31.43%
| Z-Score (Absolute) |  | 0.67 | 1.75 | 361.44 | 0 | 1.89 | 0 | 2.02 | 0.4 | 0.13 | 0.4 | 0.67 | 0.81 | 0 | 1.48 | 0.15
| Analytical results for Th in OREAS 238b (Certified Value 11.9 ppm) | [25 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | 11.695 | 12.2969 | 12.6 | 12.97 | 11.5 | 15.2 | 12.9 | 11.9 | 11.5 | 11.9 | 12.3 | 11.5 | 12.1691 | 12.51 | 12.1 | 10.18 | 15.5 | 12.6 | 8.4613 | 11.2316 |  |  |  |  | 1
| 1 | 2 | 11.819 | 12.6374 | 12.2 | 12.91 | 10.3 | 15.1 | 12.7 | 12.2 | 11.6 | 11.7 | 12.2 | 11.2 | 12.0155 | 11.4 | 11.3 | 11.6 | 15.7 | 12.5 | 8.1943 | 11.2621 |  |  |  |  | 31
| 1 | 3 | 11.64 | 12.6199 | 12.4 | 13.14 | 11.1 | 15.1 | 12.5 | 11.6 | 11.5 | 11.6 | 11.9 | 11.4 | 12.6639 | 11.92 | 12.3 | 10.15 | 16.4 | 12.6 | 7.8884 | 11.4686 |  |  |  |  | 16
| 1 | 4 | 11.584 | 12.5369 | 11.8 | 12.9 | 11.1 | 15.2 | 12.7 | 11.8 | 11.6 | 11.6 | 12.1 | 12 | 12.8296 | 11.79 | 11.2 | 10.18 | 15.9 | 12.1 | 9.6446 | 10.616 |  |  |  |  | 11.9141
| 1 | 5 | 11.827 | 12.5731 | 13 | 13.11 | 11.4 | 14.1 | 12.7 | 12.1 | 12.2 | 12.1 | 12.3 | 11.5 | 12.9066 | 11.43 | 11.7 | 9.84 | 16.6 | 12 | 9.3838 | 10.8931 |  |  |  |  | 111
| 1 | 6 | 11.38 | 12.652 | 12.7 | 13.1 | 12.4 | 15.3 | 12.7 | 12.2 | 11.4 | 11.8 | 12.2 | 10.9 | 13.2994 | 12.32 | 11.4 | 8.97 | 16.1 | 12.2 | 7.9292 | 10.9648
| Mean |  | 11.6575 | 12.5527 | 12.45 | 13.0217 | 11.3 | 15 | 12.7 | 11.9667 | 11.6333 | 11.7833 | 12.1667 | 11.4167 | 12.6474 | 11.895 | 11.6667 | 10.1533 | 16.0333 | 12.3333 | 8.5836 | 11.0727
| Median |  | 11.6675 | 12.5965 | 12.5 | 13.035 | 11.25 | 15.15 | 12.7 | 12 | 11.55 | 11.75 | 12.2 | 11.45 | 12.7467 | 11.855 | 11.55 | 10.165 | 16 | 12.35 | 8.3278 | 11.0982
| Std Dev. |  | 0.16669 | 0.13239 | 0.41833 | 0.10759 | 0.68411 | 0.44721 | 0.12649 | 0.24221 | 0.28752 | 0.19408 | 0.15055 | 0.3656 | 0.48038 | 0.45426 | 0.45019 | 0.84767 | 0.41793 | 0.26583 | 0.7542 | 0.3067
| Rel.Std.Dev. |  | 1.43% | 1.055% | 3.36% | 0.8263% | 6.054% | 2.981% | 0.996% | 2.024% | 2.472% | 1.647% | 1.237% | 3.202% | 3.798% | 3.819% | 3.859% | 8.349% | 2.607% | 2.155% | 8.787% | 2.77%
| PDM3 |  | -2.154% | 5.36% | 4.498% | 9.296% | -5.155% | 25.9% | 6.596% | 0.4409% | -2.357% | -1.098% | 2.12% | -4.175% | 6.154% | -0.1606% | -2.077% | -14.78% | 34.57% | 3.519% | -27.95% | -7.062%
| Z-Score (Absolute) |  | 0.32 | 0.73 | 0.61 | 1.29 | 0.75 | 3.63 | 0.91 | 0.04 | 0.35 | 0.17 | 0.28 | 0.61 | 0.85 | 0.04 | 0.31 | 2.1 | 4.85 | 0.48 | 3.96 | 1.01
| Analytical results for Ti in OREAS 238b (Certified Value 0.143 wt.%) | [25 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 0.1319 | 0.13 | 0.12867 | 0.1523 | 0.15 | 0.176 | 0.134 | 0.143 | 0.153 | 0.146 | 0.146 | 0.135 | 0.15616 | 0.144 | 0.13 | 0.13438 | 0.23 | 0.158 | 0.13707 | 0.3002 | 0.146 |  |  |  | 1
| 1 | 2 | 0.1317 | 0.12 | 0.12867 | 0.1511 | 0.14 | 0.179 | 0.133 | 0.147 | 0.154 | 0.144 | 0.148 | 0.137 | 0.15503 | 0.138 | 0.13 | 0.13418 | 0.23 | 0.152 | 0.13196 | 0.3088 | 0.14386 |  |  |  | 32
| 1 | 3 | 0.1314 | 0.12 | 0.12867 | 0.1571 | 0.14 | 0.177 | 0.132 | 0.142 | 0.16 | 0.146 | 0.149 | 0.137 | 0.15633 | 0.144 | 0.13 | 0.13588 | 0.24 | 0.154 | 0.12902 | 0.3148 | 0.14671 |  |  |  | 16
| 1 | 4 | 0.1314 | 0.11 | 0.12867 | 0.1557 | 0.14 | 0.181 | 0.135 | 0.145 | 0.158 | 0.145 | 0.147 | 0.14 | 0.15356 | 0.138 | 0.13 | 0.13608 | 0.22 | 0.151 | 0.13442 | 0.2668 | 0.14058 |  |  |  | 0.14283
| 1 | 5 | 0.1308 | 0.13 | 0.11967 | 0.1587 | 0.15 | 0.179 | 0.132 | 0.145 | 0.159 | 0.147 | 0.146 | 0.135 | 0.15341 | 0.138 | 0.12 | 0.14011 | 0.23 | 0.148 | 0.13553 | 0.2708 | 0.14341 |  |  |  | 112
| 1 | 6 | 0.1304 | 0.12 | 0.12867 | 0.158 | 0.14 | 0.181 | 0.135 | 0.144 | 0.159 | 0.144 | 0.146 | 0.133 | 0.15763 | 0.138 | 0.13 | 0.14103 | 0.24 | 0.151 | 0.13461 | 0.268 | 0.14578
| Mean |  | 0.13127 | 0.12167 | 0.12717 | 0.15548 | 0.14333 | 0.17883 | 0.1335 | 0.14433 | 0.15717 | 0.14533 | 0.147 | 0.13617 | 0.15535 | 0.14 | 0.12833 | 0.13694 | 0.23167 | 0.15233 | 0.13377 | 0.28823 | 0.14439
| Median |  | 0.1314 | 0.12 | 0.12867 | 0.1564 | 0.14 | 0.179 | 0.1335 | 0.1445 | 0.1585 | 0.1455 | 0.1465 | 0.136 | 0.15559 | 0.138 | 0.13 | 0.13598 | 0.23 | 0.1515 | 0.13451 | 0.2855 | 0.14482
| Std Dev. |  | 0.00056451 | 0.0075277 | 0.0036742 | 0.0031205 | 0.005164 | 0.0020412 | 0.0013784 | 0.0017512 | 0.0029269 | 0.0012111 | 0.0012649 | 0.0024014 | 0.0016651 | 0.0030984 | 0.0040825 | 0.0029263 | 0.0075277 | 0.0033862 | 0.002862 | 0.022112 | 0.0022647
| Rel.Std.Dev. |  | 0.43% | 6.187% | 2.889% | 2.007% | 3.603% | 1.141% | 1.033% | 1.213% | 1.862% | 0.8333% | 0.8605% | 1.764% | 1.072% | 2.213% | 3.181% | 2.137% | 3.249% | 2.223% | 2.139% | 7.672% | 1.568%
| PDM3 |  | -8.095% | -14.82% | -10.97% | 8.86% | 0.353% | 25.21% | -6.532% | 1.053% | 10.04% | 1.753% | 2.92% | -4.665% | 8.768% | -1.981% | -10.15% | -4.12% | 62.2% | 6.654% | -6.344% | 101.8% | 1.093%
| Z-Score (Absolute) |  | 0.81 | 1.41 | 1.07 | 0.69 | 0.06 | 2.15 | 0.67 | 0 | 0.8 | 0.06 | 0.17 | 0.51 | 0.69 | 0.28 | 1 | 0.46 | 5.44 | 0.5 | 0.66 | 8.96 | 0
| Analytical results for Tl in OREAS 238b (Certified Value 0.52 ppm) | [25 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.20 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.52 | 0.53 | < 10 | 0.55 | 0.44 | 0.58 | 0.55 | 0.52 | 0.5 | 0.5 | 0.55 | 0.52 | 0.57706 | 0.771 | 0.48 | < 5 | 0.5 | 0.4209 | 0.6415 | [5 empty] | 1
| 1 | 2 | 0.53 | 0.54 | < 10 | 0.56 | 0.42 | 0.58 | 0.55 | 0.53 | 0.5 | 0.48 | 0.54 | 0.52 | 0.589 | 0.757 | 0.46 | < 5 | 0.48 | 0.4155 | 0.6342 | [5 empty] | 33
| 1 | 3 | 0.52 | 0.53 | < 10 | 0.57 | 0.44 | 0.56 | 0.54 | 0.5 | 0.49 | 0.48 | 0.53 | 0.51 | 0.58012 | 0.78 | 0.49 | < 5 | 0.49 | 0.4011 | 0.6204 | [5 empty] | 16
| 1 | 4 | 0.51 | 0.5 | < 10 | 0.57 | 0.39 | 0.58 | 0.55 | 0.49 | 0.51 | 0.48 | 0.55 | 0.53 | 0.57323 | 0.764 | 0.47 | < 5 | 0.48 | 0.4616 | 0.6229 | [5 empty] | 0.51836
| 1 | 5 | 0.52 | 0.53 | < 10 | 0.55 | 0.41 | 0.55 | 0.54 | 0.52 | 0.51 | 0.47 | 0.53 | 0.5 | 0.58662 | 0.765 | 0.45 | < 5 | 0.48 | 0.4761 | 0.602 | [5 empty] | 113
| 1 | 6 | 0.52 | 0.52 | < 10 | 0.56 | 0.45 | 0.57 | 0.56 | 0.51 | 0.49 | 0.48 | 0.55 | 0.52 | 0.57737 | 0.771 | 0.47 | < 5 | 0.48 | 0.4094 | 0.614
| Mean |  | 0.52 | 0.525 |  | 0.56 | 0.425 | 0.57 | 0.54833 | 0.51167 | 0.5 | 0.48167 | 0.54167 | 0.51667 | 0.58056 | 0.768 | 0.47 |  | 0.485 | 0.43077 | 0.6225
| Median |  | 0.52 | 0.53 |  | 0.56 | 0.43 | 0.575 | 0.55 | 0.515 | 0.5 | 0.48 | 0.545 | 0.52 | 0.57874 | 0.768 | 0.47 |  | 0.48 | 0.4182 | 0.62165
| Std Dev. |  | 0.0063246 | 0.013784 |  | 0.0089443 | 0.022583 | 0.012649 | 0.0075277 | 0.01472 | 0.0089443 | 0.0098319 | 0.0098319 | 0.010328 | 0.0060706 | 0.0078486 | 0.014142 |  | 0.0083666 | 0.030571 | 0.014106
| Rel.Std.Dev. |  | 1.216% | 2.626% |  | 1.597% | 5.314% | 2.219% | 1.373% | 2.877% | 1.789% | 2.041% | 1.815% | 1.999% | 1.046% | 1.022% | 3.009% |  | 1.725% | 7.097% | 2.266%
| PDM3 |  | 0.3155% | 1.28% |  | 8.032% | -18.01% | 9.961% | 5.781% | -1.292% | -3.543% | -7.08% | 4.495% | -0.3275% | 12% | 48.16% | -9.33% |  | -6.436% | -16.9% | 20.09%
| Z-Score (Absolute) |  | 0.08 | 0 | 69.64 | 0.54 | 1.56 | 0.7 | 0.36 | 0.21 | 0.39 | 0.67 | 0.26 | 0.13 | 0.86 | 3.78 | 0.86 | 30.73 | 0.62 | 1.47 | 1.52
| Analytical results for Tm in OREAS 238b (Indicative Value 0.11 ppm) | [25 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | [9 empty] | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.30 | [9 empty] | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS | [9 empty] | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g | [9 empty] | 2
| 1 | 1 | 1.01: 0.1253; 1.06: 0.1; 1.07: 0.121; 1.08: 0.099 | [9 empty] | 1
| 1 | 2 | 1.01: 0.13148; 1.06: 0.1; 1.07: 0.11; 1.08: 0.1007 | [9 empty] | 13
| 1 | 3 | 1.01: 0.13419; 1.06: 0.1; 1.07: 0.117; 1.08: 0.0936 | [9 empty] | 16
| 1 | 4 | 1.01: 0.12064; 1.06: 0.1; 1.07: 0.115; 1.08: 0.111 | [9 empty] | 0.11163
| 1 | 5 | 1.01: 0.12688; 1.06: 0.1; 1.07: 0.112; 1.08: 0.1137 | [9 empty] | 19
| 1 | 6 | 1.01: 0.12683; 1.06: 0.1; 1.07: 0.124; 1.08: 0.0968
| Mean |  | 1.01: 0.12755; 1.06: 0.1; 1.07: 0.1165; 1.08: 0.10247
| Median |  | 1.01: 0.12685; 1.06: 0.1; 1.07: 0.116; 1.08: 0.09985
| Std Dev. |  | 1.01: 0.0047601; 1.06: 1.5202e-17; 1.07: 0.0053198; 1.08: 0.0080607
| Rel.Std.Dev. |  | 1.01: 3.732%; 1.06: 1.52e-14%; 1.07: 4.566%; 1.08: 7.867%
| PDM3 |  | 1.01: 14.27%; 1.06: -10.42%; 1.07: 4.362%; 1.08: -8.209%
| Z-Score (Absolute) |  | 1.01: 1.48; 1.06: 0.78; 1.07: 0.57; 1.08: 0.57
| Analytical results for U in OREAS 238b (Certified Value 1.46 ppm) | [25 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 1.4 | 1.514 | 1.5 | 1.61 | 1.42 | 1.97 | 1.5 | 1.36 | 1.42 | 1.34 | 1.54 | 1.38 | 1.5342 | 1 | 1.598 | 1.39 | 1.69 | 1.1671 | 3.5377 | [5 empty] | 1
| 1 | 2 | 1.41 | 1.612 | 1.5 | 1.61 | 1.32 | 1.98 | 1.5 | 1.39 | 1.35 | 1.33 | 1.53 | 1.43 | 1.6383 | 1 | 1.661 | 1.37 | 1.6 | 1.1168 | 3.7216 | [5 empty] | 35
| 1 | 3 | 1.409 | 1.618 | 1.4 | 1.64 | 1.29 | 1.93 | 1.5 | 1.34 | 1.36 | 1.29 | 1.51 | 1.41 | 1.5257 | 1 | 1.678 | 1.43 | 1.63 | 1.0931 | 3.5706 | [5 empty] | 16
| 1 | 4 | 1.37 | 1.4811 | 1.4 | 1.65 | 1.29 | 1.93 | 1.5 | 1.34 | 1.41 | 1.3 | 1.54 | 1.51 | 1.6221 | 1 | 1.613 | 1.4 | 1.58 | 1.2728 | 3.6513 | [5 empty] | 1.457
| 1 | 5 | 1.408 | 1.3544 | 1.5 | 1.66 | 1.29 | 1.88 | 1.5 | 1.37 | 1.45 | 1.36 | 1.53 | 1.35 | 1.4928 | 1 | 1.611 | 1.35 | 1.54 | 1.2833 | 3.697 | [5 empty] | 114
| 1 | 6 | 1.369 | 1.6403 | 1.5 | 1.64 | 1.41 | 1.93 | 1.6 | 1.41 | 1.4 | 1.3 | 1.54 | 1.4 | 1.6237 | 1 | 1.742 | 1.41 | 1.58 | 1.1249 | 3.5956
| Mean |  | 1.3943 | 1.5366 | 1.4667 | 1.635 | 1.3367 | 1.9367 | 1.5167 | 1.3683 | 1.3983 | 1.32 | 1.5317 | 1.4133 | 1.5728 | 1 | 1.6505 | 1.3917 | 1.6033 | 1.1763 | 3.629
| Median |  | 1.404 | 1.563 | 1.5 | 1.64 | 1.305 | 1.93 | 1.5 | 1.365 | 1.405 | 1.315 | 1.535 | 1.405 | 1.5782 | 1 | 1.637 | 1.395 | 1.59 | 1.146 | 3.6235
| Std Dev. |  | 0.019562 | 0.10938 | 0.05164 | 0.020736 | 0.06186 | 0.03559 | 0.040825 | 0.027869 | 0.037639 | 0.027568 | 0.01169 | 0.05465 | 0.062328 | 0 | 0.054702 | 0.028577 | 0.05164 | 0.082402 | 0.072881
| Rel.Std.Dev. |  | 1.403% | 7.118% | 3.521% | 1.268% | 4.628% | 1.838% | 2.692% | 2.037% | 2.692% | 2.088% | 0.7633% | 3.867% | 3.963% | 0% | 3.314% | 2.053% | 3.221% | 7.005% | 2.008%
| PDM3 |  | -4.302% | 5.464% | 0.6623% | 12.22% | -8.26% | 32.92% | 4.094% | -6.087% | -4.028% | -9.404% | 5.123% | -2.998% | 7.947% | -31.37% | 13.28% | -4.485% | 10.04% | -19.26% | 149.1%
| Z-Score (Absolute) |  | 0.62 | 0.29 | 0.16 | 0.91 | 0.99 | 2.84 | 0.16 | 0.79 | 0.59 | 1.09 | 0.25 | 0.5 | 0.52 | NA | 1.01 | 0.64 | 0.71 | 2.01 | 13.62
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 238b (Certified Value 69 ppm) | [25 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  | 1
| 1 | 1 | 71.2 | 73.28 | 60 | 67.3 | 69 | 69 | 69 | 72 | 67 | 67 | 70 | 66 | 68 | 67.3563 | 73 | 67 | 73.3 | 68.3 | 64 | 64.2 | 96.8081 | 70.7508 |  |  | 1
| 1 | 2 | 70.8 | 74.49 | 59 | 67.4 | 66 | 68 | 70 | 72 | 68 | 69 | 68 | 66 | 69 | 68.666 | 71 | 64 | 66.53 | 67.8 | 63 | 62.8 | 97.247 | 68.8932 |  |  | 36
| 1 | 3 | 72.6 | 76.84 | 59 | 67.1 | 66 | 68 | 70 | 72 | 66 | 69 | 69 | 67 | 70 | 70.6748 | 77 | 68 | 68.25 | 68.3 | 63 | 61.2 | 96.899 | 71.9477 |  |  | 16
| 1 | 4 | 70.4 | 74.05 | 57 | 67.22 | 66 | 67 | 71 | 73 | 67 | 69 | 69 | 66 | 70 | 72.2171 | 73 | 67 | 68.94 | 67.6 | 61 | 62.8 | 95.2915 | 68.974 |  |  | 68.5636
| 1 | 5 | 70.2 | 73.73 | 60 | 61.6467 | 64 | 70 | 71 | 72 | 68 | 70 | 69 | 66 | 68 | 69.5393 | 75 | 64 | 69.29 | 68.2 | 61 | 62.9 | 95.3678 | 68.9719 |  |  | 115
| 1 | 6 | 68.7 | 74.67 | 58 | 67.9533 | 69 | 69 | 71 | 74 | 68 | 69 | 69 | 66 | 67 | 68.8414 | 75 | 66 | 68.95 | 68.6 | 62 | 62.7 | 93.5655 | 70.5171
| Mean |  | 70.65 | 74.51 | 58.8333 | 66.4367 | 66.6667 | 68.5 | 70.3333 | 72.5 | 67.3333 | 68.8333 | 69 | 66.1667 | 68.6667 | 69.5491 | 74 | 66 | 69.21 | 68.1333 | 62.3333 | 62.7667 | 95.8632 | 70.0091
| Median |  | 70.6 | 74.27 | 59 | 67.26 | 66 | 68.5 | 70.5 | 72 | 67.5 | 69 | 69 | 66 | 68.5 | 69.1903 | 74 | 66.5 | 68.945 | 68.25 | 62.5 | 62.8 | 96.088 | 69.7455
| Std Dev. |  | 1.2802 | 1.248 | 1.169 | 2.3652 | 1.9664 | 1.0488 | 0.8165 | 0.83666 | 0.8165 | 0.98319 | 0.63246 | 0.40825 | 1.2111 | 1.7015 | 2.0976 | 1.6733 | 2.2351 | 0.36697 | 1.2111 | 0.95219 | 1.3952 | 1.2616
| Rel.Std.Dev. |  | 1.812% | 1.675% | 1.987% | 3.56% | 2.95% | 1.531% | 1.161% | 1.154% | 1.213% | 1.428% | 0.9166% | 0.617% | 1.764% | 2.446% | 2.835% | 2.535% | 3.229% | 0.5386% | 1.943% | 1.517% | 1.455% | 1.802%
| PDM3 |  | 3.043% | 8.673% | -14.19% | -3.102% | -2.767% | -0.09278% | 2.581% | 5.741% | -1.794% | 0.3934% | 0.6365% | -3.496% | 0.1503% | 1.437% | 7.929% | -3.739% | 0.9428% | -0.6276% | -9.087% | -8.455% | 39.82% | 2.108%
| Z-Score (Absolute) |  | 0.64 | 1.95 | 3.36 | 0.78 | 0.71 | 0.08 | 0.54 | 1.27 | 0.48 | 0.03 | 0.08 | 0.87 | 0.03 | 0.27 | 1.78 | 0.93 | 0.16 | 0.21 | 2.17 | 2.03 | 9.18 | 0.43
| Analytical results for W in OREAS 238b (Certified Value 0.52 ppm) | [25 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 0.61 | 0.4 | 6.804 | 0.7 | 0.36 | 0.5 | 0.6 | 0.62 | 0.65 | 0.54 | 0.61 | 0.48 | < 1 | < 1 | 0.5 | < 5 | 0.47 | 0.4633 | [6 empty] | 1
| 1 | 2 | 0.59 | 0.5 | 8.616 | 0.6 | 0.31 | 0.5 | 0.5 | 0.61 | 0.73 | 0.49 | 0.59 | 0.47 | < 1 | < 1 | 0.5 | < 5 | 0.48 | 0.4403 | [6 empty] | 37
| 1 | 3 | 0.58 | 0.4 | 6.672 | 0.6 | 0.26 | 0.5 | 0.5 | 0.61 | 0.63 | 0.51 | 0.61 | 0.45 | < 1 | < 1 | 0.5 | < 5 | 0.46 | 0.5 | [6 empty] | 16
| 1 | 4 | 0.55 | 0.4 | 9.84 | 0.6 | 0.33 | 0.4 | 0.5 | 0.56 | 0.71 | 0.62 | 0.6 | 0.44 | < 1 | < 1 | 0.4 | < 5 | 0.43 | 0.4695 | [6 empty] | 0.52
| 1 | 5 | 0.55 | 0.5 | 8.924 | 0.7 | 0.37 | 0.5 | 0.5 | 0.59 | 0.61 | 0.51 | 0.56 | 0.47 | < 1 | < 1 | 0.4 | < 5 | 0.44 | 0.5258 | [6 empty] | 116
| 1 | 6 | 0.57 | 0.5 | 7.328 | 0.7 | 0.37 | 0.5 | 0.6 | 0.6 | 0.62 | 0.5 | 0.62 | 0.44 | < 1 | < 1 | 0.5 | < 5 | 0.44 | 0.571
| Mean |  | 0.575 | 0.45 | 8.0307 | 0.65 | 0.33333 | 0.48333 | 0.53333 | 0.59833 | 0.65833 | 0.52833 | 0.59833 | 0.45833 |  |  | 0.46667 |  | 0.45333 | 0.49498
| Median |  | 0.575 | 0.45 | 7.972 | 0.65 | 0.345 | 0.5 | 0.5 | 0.605 | 0.64 | 0.51 | 0.605 | 0.46 |  |  | 0.5 |  | 0.45 | 0.48475
| Std Dev. |  | 0.023452 | 0.054772 | 1.2852 | 0.054772 | 0.043205 | 0.040825 | 0.05164 | 0.02137 | 0.049967 | 0.047924 | 0.02137 | 0.017224 |  |  | 0.05164 |  | 0.019664 | 0.047743
| Rel.Std.Dev. |  | 4.079% | 12.17% | 16% | 8.427% | 12.96% | 8.447% | 9.682% | 3.572% | 7.59% | 9.071% | 3.572% | 3.758% |  |  | 11.07% |  | 4.338% | 9.645%
| PDM3 |  | 10.58% | -13.46% | 1444% | 25% | -35.9% | -7.051% | 2.564% | 15.06% | 26.6% | 1.603% | 15.06% | -11.86% |  |  | -10.26% |  | -12.82% | -4.811%
| Z-Score (Absolute) |  | 0.45 | NA | 72.27 | NA | 1.88 | NA | NA | 0.67 | 1.25 | 0 | 0.67 | 0.67 | 0.27 | 0.27 | NA | 18.99 | 0.72 | 0.32
Results from laboratories 1.06, 1.08, 1.10, 1.11 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Y in OREAS 238b (Certified Value 9 ppm) | [25 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 8.94 | 8.5049 | 8.17 | 8.7833 | 9.79 | 8 | 9.53 | 9 | 9.13 | 9.17 | 9.05 | 9.23 | 9.7391 | 9 | 8.502 | 8.59 | 9.41 | 7.9135 | 10.3011 | 9.2025 |  |  |  |  | 1
| 1 | 2 | 9.1 | 8.8251 | 7.98 | 8.7633 | 9.87 | 7 | 9.48 | 8.9 | 9.31 | 9.01 | 8.71 | 9.35 | 9.5368 | 9 | 7.801 | 8 | 8.9 | 7.543 | 10.5872 | 9.12 |  |  |  |  | 38
| 1 | 3 | 9.21 | 8.8292 | 7.87 | 8.7467 | 10.01 | 7 | 9.67 | 8.8 | 9 | 9.63 | 8.92 | 9.15 | 9.8126 | 9 | 8.231 | 8.99 | 8.99 | 7.543 | 10.8022 | 8.9479 |  |  |  |  | 16
| 1 | 4 | 9.04 | 8.2753 | 7.48 | 8.8 | 9.89 | 8 | 9.55 | 9.1 | 9.11 | 9.8 | 8.65 | 9.34 | 9.7715 | 9 | 8.174 | 8.62 | 8.99 | 7.638 | 9.9808 | 8.7719 |  |  |  |  | 8.9951
| 1 | 5 | 9.1 | 8.3962 | 7.86 | 8.0867 | 10 | 8 | 9.73 | 8.9 | 9.35 | 9.22 | 9.05 | 9.3 | 9.7225 | 9 | 7.867 | 8.22 | 8.64 | 7.695 | 9.8368 | 8.8802 |  |  |  |  | 117
| 1 | 6 | 9.07 | 8.8635 | 7.93 | 8.9033 | 10.04 | 8 | 9.79 | 9.2 | 9.45 | 9.09 | 8.58 | 8.89 | 9.6186 | 9 | 8.667 | 9.16 | 8.9 | 7.6 | 9.8709 | 8.9506
| Mean |  | 9.0767 | 8.6157 | 7.8817 | 8.6806 | 9.9333 | 7.6667 | 9.625 | 8.9833 | 9.225 | 9.32 | 8.8267 | 9.21 | 9.7002 | 9 | 8.207 | 8.5967 | 8.9717 | 7.6554 | 10.2298 | 8.9788
| Median |  | 9.085 | 8.665 | 7.9 | 8.7733 | 9.945 | 8 | 9.61 | 8.95 | 9.22 | 9.195 | 8.815 | 9.265 | 9.7308 | 9 | 8.2025 | 8.605 | 8.945 | 7.619 | 10.141 | 8.9493
| Std Dev. |  | 0.088242 | 0.25581 | 0.22675 | 0.2961 | 0.098116 | 0.5164 | 0.12325 | 0.1472 | 0.17108 | 0.31875 | 0.20695 | 0.17378 | 0.10303 | 0 | 0.3408 | 0.44058 | 0.25023 | 0.13916 | 0.40143 | 0.15766
| Rel.Std.Dev. |  | 0.9722% | 2.969% | 2.877% | 3.411% | 0.9877% | 6.736% | 1.28% | 1.639% | 1.855% | 3.42% | 2.345% | 1.887% | 1.062% | 0% | 4.153% | 5.125% | 2.789% | 1.818% | 3.924% | 1.756%
| PDM3 |  | 0.9069% | -4.218% | -12.38% | -3.497% | 10.43% | -14.77% | 7.003% | -0.1307% | 2.556% | 3.612% | -1.872% | 2.389% | 7.838% | 0.05463% | -8.761% | -4.429% | -0.2604% | -14.89% | 13.73% | -0.1806%
| Z-Score (Absolute) |  | 0.18 | 0.67 | 2.04 | 0.55 | 1.77 | 2.44 | 1.2 | 0.01 | 0.46 | 0.63 | 0.28 | 0.43 | 1.34 | NA | 1.43 | 0.71 | 0.01 | 2.46 | 2.32 | 0
Results from laboratory 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 238b (Certified Value 0.87 ppm) | [25 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab | [9 empty] | 1
| No | No | 1.01: 1.03; 1.06: 1.06; 1.07: 1.20; 1.08: 1.21; 1.09: 1.30; 1.10: 1.32 | [9 empty] | ppm
|  |  | 1.01: AR*MS; 1.06: AR*OES/MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*OES; 1.10: AR*OES | [9 empty] | 2
|  |  | 1.01: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.5g; 1.10: 0.5g | [9 empty] | 2
| 1 | 1 | 1.01: 0.85367; 1.06: 0.8; 1.07: 0.76; 1.08: 0.8; 1.09: 0.97; 1.10: 1.1406 | [9 empty] | 1
| 1 | 2 | 1.01: 0.83562; 1.06: 0.8; 1.07: 0.695; 1.08: 0.7; 1.09: 0.96; 1.10: 1.1702 | [9 empty] | 39
| 1 | 3 | 1.01: 0.84344; 1.06: 0.8; 1.07: 0.736; 1.08: 0.8; 1.09: 0.92; 1.10: 1.1999 | [9 empty] | 16
| 1 | 4 | 1.01: 0.87219; 1.06: 0.7; 1.07: 0.727; 1.08: 0.7; 1.09: 0.95; 1.10: 1.0227 | [9 empty] | 0.86718
| 1 | 5 | 1.01: 0.89463; 1.06: 0.8; 1.07: 0.693; 1.08: 0.8; 1.09: 1; 1.10: 1.0815 | [9 empty] | 118
| 1 | 6 | 1.01: 0.86555; 1.06: 0.8; 1.07: 0.786; 1.08: 0.7; 1.09: 0.96; 1.10: 1.0813
| Mean |  | 1.01: 0.86085; 1.06: 0.78333; 1.07: 0.73283; 1.08: 0.75; 1.09: 0.96; 1.10: 1.116
| Median |  | 1.01: 0.85961; 1.06: 0.8; 1.07: 0.7315; 1.08: 0.75; 1.09: 0.96; 1.10: 1.1111
| Std Dev. |  | 1.01: 0.021369; 1.06: 0.040825; 1.07: 0.036406; 1.08: 0.054772; 1.09: 0.026077; 1.10: 0.06585
| Rel.Std.Dev. |  | 1.01: 2.482%; 1.06: 5.212%; 1.07: 4.968%; 1.08: 7.303%; 1.09: 2.716%; 1.10: 5.9%
| PDM3 |  | 1.01: -0.7292%; 1.06: -9.668%; 1.07: -15.49%; 1.08: -13.51%; 1.09: 10.7%; 1.10: 28.7%
| Z-Score (Absolute) |  | 1.01: 0.32; 1.06: 0.32; 1.07: 0.75; 1.08: 0.6; 1.09: 1.15; 1.10: 2.46
| Analytical results for Zn in OREAS 238b (Certified Value 87 ppm) | [25 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.25 | 1.27 | 1.30 | 1.31 | 1.32 | 1.34 | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.2g | 01g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 1
| 1 | 1 | 81.5 | 86.735 | 87.3 | 83.7133 | 85 | 88 | 88 | 88 | 88 | 85 | 93 | 86 | 87.1 | 90.1683 | 90 | 90.65 | 81 | 90.91 | 85 | 80 | 101.7184 | 93 | 82.4808 | 90.6436 | 1
| 1 | 2 | 81.8 | 89.87 | 89 | 81.52 | 87 | 82 | 90 | 86 | 87 | 84 | 88 | 87 | 87.5 | 89.6392 | 87 | 89.91 | 76 | 88.49 | 85.9 | 78 | 101.1248 | 95 | 80.1993 | 87.0318 | 40
| 1 | 3 | 81.7 | 90.354 | 86.6 | 84.77 | 86 | 83 | 90 | 85 | 86 | 86 | 89 | 87 | 87.3 | 90.4976 | 90 | 90.94 | 81 | 92.59 | 86.4 | 79 | 97.9328 | 93 | 82.0554 | 91.4422 | 16
| 1 | 4 | 80.5 | 90.31 | 84.7 | 82.615 | 86 | 83 | 90 | 89 | 87 | 85 | 88 | 86 | 88.9 | 86.6569 | 86 | 90.96 | 81 | 93.22 | 82.5 | 76 | 100.0944 | 96 | 76.58 | 88.231 | 86.8375
| 1 | 5 | 81.3 | 90.717 | 87.3 | 82.495 | 84 | 88 | 90 | 86 | 87 | 85 | 88 | 87 | 85.9 | 88.4444 | 89 | 92.09 | 78 | 93.73 | 84.9 | 76 | 100.1728 | 98 | 75.5376 | 87.2611 | 119
| 1 | 6 | 81.5 | 89.331 | 85 | 82.8233 | 87 | 79 | 91 | 90 | 87 | 85 | 88 | 87 | 85.2 | 87.768 | 88 | 99.53 | 78 | 92.58 | 85.7 | 77 | 99.5232 | 98 | 75.0979 | 89.3024
| Mean |  | 81.3833 | 89.5528 | 86.65 | 82.9894 | 85.8333 | 83.8333 | 89.8333 | 87.3333 | 87 | 85 | 89 | 86.6667 | 86.9833 | 88.8624 | 88.3333 | 92.3467 | 79.1667 | 91.92 | 85.0667 | 77.6667 | 100.0944 | 95.5 | 78.6585 | 88.9853
| Median |  | 81.5 | 90.09 | 86.95 | 82.7192 | 86 | 83 | 90 | 87 | 87 | 85 | 88 | 87 | 87.2 | 89.0418 | 88.5 | 90.95 | 79.5 | 92.585 | 85.35 | 77.5 | 100.1336 | 95.5 | 78.3897 | 88.7667
| Std Dev. |  | 0.46655 | 1.4602 | 1.6059 | 1.1195 | 1.169 | 3.5449 | 0.98319 | 1.9664 | 0.63246 | 0.63246 | 2 | 0.5164 | 1.2968 | 1.498 | 1.633 | 3.5884 | 2.137 | 1.9306 | 1.3779 | 1.633 | 1.3194 | 2.2583 | 3.3245 | 1.8025
| Rel.Std.Dev. |  | 0.5733% | 1.631% | 1.853% | 1.349% | 1.362% | 4.229% | 1.094% | 2.252% | 0.727% | 0.7441% | 2.247% | 0.5958% | 1.491% | 1.686% | 1.849% | 3.886% | 2.699% | 2.1% | 1.62% | 2.103% | 1.318% | 2.365% | 4.226% | 2.026%
| PDM3 |  | -6.281% | 3.127% | -0.2159% | -4.431% | -1.156% | -3.46% | 3.45% | 0.571% | 0.1871% | -2.116% | 2.49% | -0.1967% | 0.168% | 2.332% | 1.723% | 6.344% | -8.834% | 5.853% | -2.039% | -10.56% | 15.27% | 9.976% | -9.419% | 2.473%
| Z-Score (Absolute) |  | 1.66 | 0.76 | 0.1 | 1.18 | 0.34 | 0.93 | 0.84 | 0.1 | 0 | 0.59 | 0.59 | 0.1 | 0 | 0.55 | 0.4 | 1.58 | 2.31 | 1.45 | 0.57 | 2.75 | 3.87 | 2.51 | 2.46 | 0.59
| Analytical results for Zr in OREAS 238b (Certified Value 21.3 ppm) | [25 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.19 | 1.21 | 1.25 | 1.27 | 1.30 | 1.32 | 1.34 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*MS |  |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.2g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.5g | 0.5g | 0.5g | 0.5g |  |  |  |  | 1
| 1 | 1 | 20.28 | 31.43 | 5.9 | 22.8782 | 21.3 | 12 | 26.2 | 19.9 | 20.8 | 23.6 | 22 | 22 | 24.1886 | 14 | 17.7 | 23.7 | 24.2 | 15.76 | 96.0157 | 20.6002 |  |  |  |  | 1
| 1 | 2 | 20.54 | 32.03 | 6.1 | 22.8194 | 21.3 | 11.7 | 26.8 | 19.5 | 21.2 | 20.9 | 21.4 | 21.8 | 24.1619 | 12 | 17 | 22.1 | 23.4 | 15.7 | 96.2268 | 20.2533 |  |  |  |  | 41
| 1 | 3 | 20.98 | 33.09 | 6.7 | 22.7491 | 21.5 | 12.2 | 25.6 | 19 | 20.7 | 22.4 | 21.2 | 20.8 | 23.5901 | 13 | 17.3 | 24.1 | 23.9 | 15.4 | 96.7192 | 19.9013 |  |  |  |  | 16
| 1 | 4 | 20.8 | 32.4 | 5.9 | 22.915 | 21.3 | 11.2 | 25.9 | 20 | 20.8 | 21.7 | 20.4 | 21.7 | 23.8793 | 13 | 16.9 | 22.7 | 22.8 | 15.59 | 93.0662 | 18.8373 |  |  |  |  | 21.2593
| 1 | 5 | 20.92 | 31.69 | 9.9 | 21.1796 | 21.2 | 10.3 | 25.3 | 19.4 | 21.3 | 20.9 | 22.3 | 21.4 | 23.793 | 13 | 16.4 | 25 | 22.1 | 15.45 | 94.0386 | 19.3616 |  |  |  |  | 120
| 1 | 6 | 20.33 | 32.15 | 10.7 | 23.3208 | 21.5 | 12.8 | 25.9 | 19.9 | 21.5 | 21.2 | 20.5 | 20.9 | 24.1953 | 13 | 17.5 | 22.1 | 22.8 | 15.22 | 94.1214 | 19.3848
| Mean |  | 20.6417 | 32.1317 | 7.5333 | 22.6437 | 21.35 | 11.7 | 25.95 | 19.6167 | 21.05 | 21.7833 | 21.3 | 21.4333 | 23.968 | 13 | 17.1333 | 23.2833 | 23.2 | 15.52 | 95.0313 | 19.7231
| Median |  | 20.67 | 32.09 | 6.4 | 22.8488 | 21.3 | 11.85 | 25.9 | 19.7 | 21 | 21.45 | 21.3 | 21.55 | 24.0206 | 13 | 17.15 | 23.2 | 23.1 | 15.52 | 95.0686 | 19.643
| Std Dev. |  | 0.30176 | 0.58112 | 2.1778 | 0.74466 | 0.12247 | 0.86718 | 0.51672 | 0.38687 | 0.32711 | 1.0572 | 0.76942 | 0.49261 | 0.25266 | 0.63246 | 0.46762 | 1.1771 | 0.7823 | 0.20209 | 1.478 | 0.64992
| Rel.Std.Dev. |  | 1.462% | 1.809% | 28.91% | 3.289% | 0.5737% | 7.412% | 1.991% | 1.972% | 1.554% | 4.853% | 3.612% | 2.298% | 1.054% | 4.865% | 2.729% | 5.056% | 3.372% | 1.302% | 1.555% | 3.295%
| PDM3 |  | -2.905% | 51.14% | -64.56% | 6.512% | 0.4267% | -44.97% | 22.06% | -7.727% | -0.9844% | 2.465% | 0.1915% | 0.8187% | 12.74% | -38.85% | -19.41% | 9.521% | 9.129% | -27% | 347% | -7.226%
| Z-Score (Absolute) |  | 0.24 | 3.8 | 4.85 | 0.46 | 0.01 | 3.39 | 1.63 | 0.6 | 0.1 | 0.16 | 0.01 | 0.04 | 0.93 | 2.93 | 1.47 | 0.69 | 0.66 | 2.04 | 25.93 | 0.56
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C201: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C457: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C493: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C511: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C529: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C547: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C565: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C583: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C908: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C980: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C998: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1016: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1124: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 238b (Indicative Value 14.65 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.66 | 1
| 1 | 2 | 14.63 | 16
| Mean |  | 14.645 | 16
| Median |  | 14.645 | 14.645
| Std Dev. |  | 0.021213 | 22
| Rel.Std.Dev. |  | 0.1448%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 238b (Indicative Value 1.2 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.19 | 17
| Mean |  | 1.195 | 16
| Median |  | 1.195 | 1.195
| Std Dev. |  | 0.0070711 | 23
| Rel.Std.Dev. |  | 0.5917%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 238b (Indicative Value 5.88 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.88 | 1
| 1 | 2 | 5.88 | 18
| Mean |  | 5.88 | 16
| Median |  | 5.88 | 5.88
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 238b (Indicative Value 3.34 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.34 | 1
| 1 | 2 | 3.34 | 19
| Mean |  | 3.34 | 16
| Median |  | 3.34 | 3.34
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 238b (Indicative Value 2.93 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.93 | 1
| 1 | 2 | 2.92 | 16
| Mean |  | 2.925 | 16
| Median |  | 2.925 | 2.925
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 0.2417%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 238b (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 17
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 238b (Indicative Value 0.945 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.94 | 1
| 1 | 2 | 0.95 | 18
| Mean |  | 0.945 | 16
| Median |  | 0.945 | 0.945
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 238b (Indicative Value 0.151 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.151 | 1
| 1 | 2 | 0.15 | 19
| Mean |  | 0.1505 | 16
| Median |  | 0.1505 | 0.1505
| Std Dev. |  | 0.00070711 | 25
| Rel.Std.Dev. |  | 0.4698%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 238b (Indicative Value 67.39 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 67.35 | 1
| 1 | 2 | 67.42 | 16
| Mean |  | 67.385 | 16
| Median |  | 67.385 | 67.385
| Std Dev. |  | 0.049497 | 22
| Rel.Std.Dev. |  | 0.07345%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 238b (Indicative Value 0.436 wt.%) |  |  | Indicative
SO3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.432 | 1
| 1 | 2 | 0.439 | 17
| Mean |  | 0.4355 | 16
| Median |  | 0.4355 | 0.4355
| Std Dev. |  | 0.0049497 | 23
| Rel.Std.Dev. |  | 1.137%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 238b (Indicative Value 0.82 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.82 | 1
| 1 | 2 | 0.82 | 18
| Mean |  | 0.82 | 16
| Median |  | 0.82 | 0.82
| Std Dev. |  | 0 | 24
| Rel.Std.Dev. |  | 0%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 238b (Indicative Value 2.5 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 2.51 | 1
| 1 | 2 | 2.49 | 21
| Mean |  | 2.5 | 16
| Median |  | 2.5 | 2.5
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 0.5657%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 238b (Indicative Value 0.105 wt.%) |  |  |  | Indicative
C | Round | Replicate | BV Geo |  | 1
| No | No | Perth |  | wt.%
|  |  | IRC |  | 3
3
| 1 | 1 | 0.11 |  | 1
| 1 | 2 | 0.1 |  | 23
| Mean |  | 0.105 |  | 16
| Median |  | 0.105 |  | 0.105
| Std Dev. |  | 0.0070711 |  | 29
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 238b (Indicative Value 0.167 wt.%) |  |  |  | Indicative
S | Round | Replicate | Inspectorate | BV Geo | 1
| No | No | Gandhidham | Perth | wt.%
|  |  | IRC | IRC | 3
3
| 1 | 1 | 0.16 | 0.17 | 1
| 1 | 2 | 0.17 | 0.16 | 23
| 1 | 3 | 0.17 |  | 16
| 1 | 4 | 0.17 |  | 0.16667
| 1 | 5 | 0.17 |  | 29
| 1 | 6 | 0.17
| Mean |  | 0.16833 | 0.165
| Median |  | 0.17 | 0.165
| Std Dev. |  | 0.0040825 | 0.0070711
| Rel.Std.Dev. |  | 2.425% | 4.285%
| PDM3 |  | 1% | -1%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 238b (Indicative Value 0.35 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.3 | 25
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.070711 | 31
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 238b (Indicative Value 698 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 689 | 1
| 1 | 2 | 706 | 26
| Mean |  | 697.5 | 16
| Median |  | 697.5 | 697.5
| Std Dev. |  | 12.0208 | 32
| Rel.Std.Dev. |  | 1.723%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 238b (Indicative Value 697 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 688 | 1
| 1 | 2 | 705 | 27
| Mean |  | 696.5 | 16
| Median |  | 696.5 | 696.5
| Std Dev. |  | 12.0208 | 33
| Rel.Std.Dev. |  | 1.726%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 238b (Indicative Value 3 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.2 | 1
| 1 | 2 | 2.8 | 28
| Mean |  | 3 | 16
| Median |  | 3 | 3
| Std Dev. |  | 0.28284 | 34
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 238b (Indicative Value 0.51 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.48 | 1
| 1 | 2 | 0.54 | 29
| Mean |  | 0.51 | 16
| Median |  | 0.51 | 0.51
| Std Dev. |  | 0.042426 | 35
| Rel.Std.Dev. |  | 8.319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 238b (Indicative Value 0.2 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 30
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 238b (Indicative Value 80 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 79.5 | 1
| 1 | 2 | 81 | 31
| Mean |  | 80.25 | 16
| Median |  | 80.25 | 80.25
| Std Dev. |  | 1.0607 | 37
| Rel.Std.Dev. |  | 1.322%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 238b (Indicative Value 18 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.7 | 1
| 1 | 2 | 18.3 | 32
| Mean |  | 18 | 16
| Median |  | 18 | 18
| Std Dev. |  | 0.42426 | 38
| Rel.Std.Dev. |  | 2.357%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 238b (Indicative Value 150 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 150 | 1
| 1 | 2 | 150 | 33
| Mean |  | 150 | 16
| Median |  | 150 | 150
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 238b (Indicative Value 9.56 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.6 | 1
| 1 | 2 | 9.51 | 34
| Mean |  | 9.555 | 16
| Median |  | 9.555 | 9.555
| Std Dev. |  | 0.06364 | 40
| Rel.Std.Dev. |  | 0.666%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 238b (Indicative Value 31 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 32 | 35
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 1.4142 | 41
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 238b (Indicative Value 5.87 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.88 | 1
| 1 | 2 | 5.85 | 9
| Mean |  | 5.865 | 16
| Median |  | 5.865 | 5.865
| Std Dev. |  | 0.021213 | 42
| Rel.Std.Dev. |  | 0.3617%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 238b (Indicative Value 3.37 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.39 | 1
| 1 | 2 | 3.34 | 10
| Mean |  | 3.365 | 16
| Median |  | 3.365 | 3.365
| Std Dev. |  | 0.035355 | 43
| Rel.Std.Dev. |  | 1.051%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 238b (Indicative Value 1.36 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.34 | 1
| 1 | 2 | 1.37 | 11
| Mean |  | 1.355 | 16
| Median |  | 1.355 | 1.355
| Std Dev. |  | 0.021213 | 44
| Rel.Std.Dev. |  | 1.566%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 238b (Indicative Value 19.8 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.8 | 1
| 1 | 2 | 19.7 | 39
| Mean |  | 19.75 | 16
| Median |  | 19.75 | 19.75
| Std Dev. |  | 0.070711 | 45
| Rel.Std.Dev. |  | 0.358%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 238b (Indicative Value 6.22 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.14 | 1
| 1 | 2 | 6.3 | 12
| Mean |  | 6.22 | 16
| Median |  | 6.22 | 6.22
| Std Dev. |  | 0.11314 | 46
| Rel.Std.Dev. |  | 1.819%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 238b (Indicative Value 1.55 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 1.5 | 5
| Mean |  | 1.55 | 16
| Median |  | 1.55 | 1.55
| Std Dev. |  | 0.070711 | 47
| Rel.Std.Dev. |  | 4.562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 238b (Indicative Value 6.64 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.55 | 1
| 1 | 2 | 6.72 | 25
| Mean |  | 6.635 | 16
| Median |  | 6.635 | 6.635
| Std Dev. |  | 0.12021 | 31
| Rel.Std.Dev. |  | 1.812%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 238b (Indicative Value 1.18 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.17 | 1
| 1 | 2 | 1.18 | 14
| Mean |  | 1.175 | 16
| Median |  | 1.175 | 1.175
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 0.6018%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 238b (Indicative Value < 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.05 | 1
| 1 | 2 | < 0.05 | 27
| Mean |  |  | 16
| Median |  |  | < 0.05
| Std Dev. |  |  | 33
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 238b (Indicative Value 41.2 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 40.8 | 1
| 1 | 2 | 41.5 | 28
| Mean |  | 41.15 | 16
| Median |  | 41.15 | 41.15
| Std Dev. |  | 0.49497 | 34
| Rel.Std.Dev. |  | 1.203%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 238b (Indicative Value 0.47 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.46 | 1
| 1 | 2 | 0.48 | 29
| Mean |  | 0.47 | 16
| Median |  | 0.47 | 0.47
| Std Dev. |  | 0.014142 | 35
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 238b (Indicative Value 0.042 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0416 | 1
| 1 | 2 | 0.0426 | 30
| Mean |  | 0.0421 | 16
| Median |  | 0.0421 | 0.0421
| Std Dev. |  | 0.00070711 | 36
| Rel.Std.Dev. |  | 1.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 238b (Indicative Value 1 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1 | 1
| 1 | 2 | 1 | 31
| Mean |  | 1 | 16
| Median |  | 1 | 1
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 238b (Indicative Value 16 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.9 | 1
| 1 | 2 | 16 | 32
| Mean |  | 15.95 | 16
| Median |  | 15.95 | 15.95
| Std Dev. |  | 0.070711 | 38
| Rel.Std.Dev. |  | 0.4433%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 238b (Indicative Value 37 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.3 | 1
| 1 | 2 | 37.6 | 9
| Mean |  | 36.95 | 16
| Median |  | 36.95 | 36.95
| Std Dev. |  | 0.91924 | 39
| Rel.Std.Dev. |  | 2.488%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 238b (Indicative Value 69 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 68 | 1
| 1 | 2 | 70 | 34
| Mean |  | 69 | 16
| Median |  | 69 | 69
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 2.05%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 238b (Indicative Value 26.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27 | 1
| 1 | 2 | 26 | 35
| Mean |  | 26.5 | 16
| Median |  | 26.5 | 26.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 2.668%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 238b (Indicative Value 9.79 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.79 | 1
| 1 | 2 | 9.79 | 11
| Mean |  | 9.79 | 16
| Median |  | 9.79 | 9.79
| Std Dev. |  | 0 | 42
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 238b (Indicative Value 162 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 162 | 1
| 1 | 2 | 162 | 37
| Mean |  | 162 | 16
| Median |  | 162 | 162
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 238b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 238b (Indicative Value 664 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 658 | 1
| 1 | 2 | 669 | 39
| Mean |  | 663.5 | 16
| Median |  | 663.5 | 663.5
| Std Dev. |  | 7.7782 | 45
| Rel.Std.Dev. |  | 1.172%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 238b (Indicative Value 14.8 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.7 | 1
| 1 | 2 | 14.8 | 40
| Mean |  | 14.75 | 16
| Median |  | 14.75 | 14.75
| Std Dev. |  | 0.070711 | 46
| Rel.Std.Dev. |  | 0.4794%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 238b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 238b (Indicative Value 7.35 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.22 | 1
| 1 | 2 | 7.47 | 10
| Mean |  | 7.345 | 16
| Median |  | 7.345 | 7.345
| Std Dev. |  | 0.17678 | 31
| Rel.Std.Dev. |  | 2.407%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 238b (Indicative Value 3.8 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 3.6 | 26
| Mean |  | 3.8 | 16
| Median |  | 3.8 | 3.8
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 238b (Indicative Value 118 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 117 | 1
| 1 | 2 | 118 | 27
| Mean |  | 117.5 | 16
| Median |  | 117.5 | 117.5
| Std Dev. |  | 0.70711 | 33
| Rel.Std.Dev. |  | 0.6018%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 238b (Indicative Value 1.22 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.23 | 28
| Mean |  | 1.215 | 16
| Median |  | 1.215 | 1.215
| Std Dev. |  | 0.021213 | 34
| Rel.Std.Dev. |  | 1.746%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 238b (Indicative Value 0.97 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.97 | 1
| 1 | 2 | 0.96 | 11
| Mean |  | 0.965 | 16
| Median |  | 0.965 | 0.965
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 238b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 5
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 238b (Indicative Value 15.4 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15.4 | 1
| 1 | 2 | 15.4 | 31
| Mean |  | 15.4 | 16
| Median |  | 15.4 | 15.4
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 238b (Indicative Value 0.501 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.502 | 32
| Mean |  | 0.501 | 16
| Median |  | 0.501 | 0.501
| Std Dev. |  | 0.0014142 | 38
| Rel.Std.Dev. |  | 0.2823%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 238b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 33
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 238b (Indicative Value 0.51 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.51 | 1
| 1 | 2 | 0.5 | 13
| Mean |  | 0.505 | 16
| Median |  | 0.505 | 0.505
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.4%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 238b (Indicative Value 3.12 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.11 | 1
| 1 | 2 | 3.13 | 35
| Mean |  | 3.12 | 16
| Median |  | 3.12 | 3.12
| Std Dev. |  | 0.014142 | 41
| Rel.Std.Dev. |  | 0.4533%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 238b (Indicative Value 113 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 112 | 1
| 1 | 2 | 114 | 36
| Mean |  | 113 | 16
| Median |  | 113 | 113
| Std Dev. |  | 1.4142 | 42
| Rel.Std.Dev. |  | 1.252%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 238b (Indicative Value 4 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4 | 1
| 1 | 2 | 4 | 37
| Mean |  | 4 | 16
| Median |  | 4 | 4
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 238b (Indicative Value 31 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31.1 | 1
| 1 | 2 | 30.9 | 38
| Mean |  | 31 | 16
| Median |  | 31 | 31
| Std Dev. |  | 0.14142 | 44
| Rel.Std.Dev. |  | 0.4562%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 238b (Indicative Value 3.22 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.25 | 1
| 1 | 2 | 3.18 | 39
| Mean |  | 3.215 | 16
| Median |  | 3.215 | 3.215
| Std Dev. |  | 0.049497 | 45
| Rel.Std.Dev. |  | 1.54%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 238b (Indicative Value 100 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 40
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 46
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 238b (Indicative Value 234 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 232 | 1
| 1 | 2 | 235 | 41
| Mean |  | 233.5 | 16
| Median |  | 233.5 | 233.5
| Std Dev. |  | 2.1213 | 47
| Rel.Std.Dev. |  | 0.9085%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.