## Performance Gates
Table 1. Certified Values and Performance Gates for OREAS 931b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
4-Acid Digestion
Ag, ppm | 16.8756 | 0.54214 | 15.7913 | 17.9599 | 15.2492 | 18.502 | 3.213% | 6.425% | 9.638% | 16.0318 | 17.7194
Al, wt.% | 5.5432 | 0.15521 | 5.2328 | 5.8536 | 5.0776 | 6.0088 | 2.8% | 5.6% | 8.4% | 5.266 | 5.8204
As, ppm | 98.5527 | 4.5426 | 89.4675 | 107.638 | 84.9248 | 112.1806 | 4.609% | 9.219% | 13.83% | 93.6251 | 103.4803
Be, ppm | 2.0818 | 0.18435 | 1.7131 | 2.4505 | 1.5287 | 2.6348 | 8.855% | 17.71% | 26.57% | 1.9777 | 2.1858
Bi, ppm | 211.3386 | 14.9804 | 181.3777 | 241.2994 | 166.3973 | 256.2799 | 7.088% | 14.18% | 21.27% | 200.7717 | 221.9055
Ca, wt.% | 0.083318 | 0.0045206 | 0.074277 | 0.092359 | 0.069756 | 0.096879 | 5.426% | 10.85% | 16.28% | 0.079152 | 0.087484
Cd, ppm | 0.64385 | 0.064894 | 0.51406 | 0.77364 | 0.44917 | 0.83853 | 10.08% | 20.16% | 30.24% | 0.61166 | 0.67604
Ce, ppm | 70.6165 | 5.0816 | 60.4533 | 80.7797 | 55.3717 | 85.8613 | 7.196% | 14.39% | 21.59% | 67.0857 | 74.1473
Co, ppm | 29.5646 | 1.4094 | 26.7458 | 32.3834 | 25.3364 | 33.7928 | 4.767% | 9.534% | 14.3% | 28.0864 | 31.0429
Cr, ppm | 72.1313 | 6.2025 | 59.7263 | 84.5362 | 53.5238 | 90.7387 | 8.599% | 17.2% | 25.8% | 68.5247 | 75.7378
Cs, ppm | 4.7453 | 0.26068 | 4.2239 | 5.2666 | 3.9633 | 5.5273 | 5.493% | 10.99% | 16.48% | 4.508 | 4.9825
Cu, wt.% | 3.9471 | 0.1243 | 3.6985 | 4.1957 | 3.5742 | 4.32 | 3.149% | 6.298% | 9.448% | 3.7498 | 4.1445
Dy, ppm | 3.0608 | 0.32 | 2.4208 | 3.7008 | 2.1008 | 4.0208 | 10.45% | 20.91% | 31.36% | 2.9078 | 3.2139
Er, ppm | 1.3743 | 0.091404 | 1.1915 | 1.5571 | 1.1 | 1.6485 | 6.651% | 13.3% | 19.95% | 1.3055 | 1.443
Eu, ppm | 1.0906 | 0.10587 | 0.8789 | 1.3024 | 0.77302 | 1.4082 | 9.707% | 19.41% | 29.12% | 1.0361 | 1.1452
Fe, wt.% | 8.552 | 0.34148 | 7.869 | 9.2349 | 7.5275 | 9.5764 | 3.993% | 7.986% | 11.98% | 8.1244 | 8.9796
Ga, ppm | 17.7095 | 1.2261 | 15.2572 | 20.1618 | 14.0311 | 21.388 | 6.924% | 13.85% | 20.77% | 16.8241 | 18.595
Gd, ppm | 4.4278 | 0.41073 | 3.6063 | 5.2492 | 3.1956 | 5.66 | 9.276% | 18.55% | 27.83% | 4.2064 | 4.6492
Ge, ppm | 0.20952 | 0.037114 | 0.1353 | 0.28375 | 0.098182 | 0.32087 | 17.71% | 35.43% | 53.14% | 0.19905 | 0.22
Hf, ppm | 1.8155 | 0.08032 | 1.6549 | 1.9762 | 1.5746 | 2.0565 | 4.424% | 8.848% | 13.27% | 1.7248 | 1.9063
Ho, ppm | 0.48863 | 0.04375 | 0.40113 | 0.57613 | 0.35738 | 0.61988 | 8.954% | 17.91% | 26.86% | 0.4642 | 0.51306
In, ppm | 3.3152 | 0.14453 | 3.0262 | 3.6043 | 2.8816 | 3.7488 | 4.36% | 8.719% | 13.08% | 3.1495 | 3.481
K, wt.% | 2.1761 | 0.082879 | 2.0104 | 2.3419 | 1.9275 | 2.4248 | 3.809% | 7.617% | 11.43% | 2.0673 | 2.285
La, ppm | 32.3429 | 3.7114 | 24.92 | 39.7657 | 21.2086 | 43.4772 | 11.48% | 22.95% | 34.43% | 30.7257 | 33.96
Li, ppm | 16.9919 | 0.6457 | 15.7005 | 18.2833 | 15.0547 | 18.929 | 3.8% | 7.6% | 11.4% | 16.1423 | 17.8414
Lu, ppm | 0.22246 | 0.016594 | 0.18928 | 0.25565 | 0.17268 | 0.27225 | 7.459% | 14.92% | 22.38% | 0.21134 | 0.23359
Mg, wt.% | 0.35396 | 0.019331 | 0.3153 | 0.39262 | 0.29597 | 0.41195 | 5.461% | 10.92% | 16.38% | 0.33626 | 0.37166
Mn, wt.% | 0.016092 | 0.00064331 | 0.014805 | 0.017379 | 0.014162 | 0.018022 | 3.998% | 7.995% | 11.99% | 0.015287 | 0.016897
Mo, ppm | 3.3832 | 0.17508 | 3.033 | 3.7333 | 2.8579 | 3.9084 | 5.175% | 10.35% | 15.52% | 3.214 | 3.5523
Na, wt.% | 0.076484 | 0.004799 | 0.066886 | 0.086082 | 0.062087 | 0.090881 | 6.275% | 12.55% | 18.82% | 0.07266 | 0.080308
Nb, ppm | 2.7246 | 0.62596 | 1.4727 | 3.9765 | 0.8467 | 4.6024 | 22.97% | 45.95% | 68.92% | 2.5883 | 2.8608
Nd, ppm | 29.8611 | 2.5201 | 24.821 | 34.9013 | 22.3009 | 37.4213 | 8.439% | 16.88% | 25.32% | 28.3681 | 31.3542
Ni, ppm | 12.5914 | 0.67895 | 11.2335 | 13.9493 | 10.5546 | 14.6283 | 5.392% | 10.78% | 16.18% | 11.9619 | 13.221
P, wt.% | 0.027157 | 0.0015324 | 0.024092 | 0.030222 | 0.02256 | 0.031754 | 5.643% | 11.29% | 16.93% | 0.025799 | 0.028515
Pb, ppm | 121.0905 | 6.5938 | 107.9029 | 134.2781 | 101.3091 | 140.8719 | 5.445% | 10.89% | 16.34% | 115.036 | 127.1451
Pr, ppm | 8.1238 | 0.78008 | 6.5636 | 9.6839 | 5.7836 | 10.464 | 9.602% | 19.2% | 28.81% | 7.7176 | 8.53
Rb, ppm | 126.6675 | 4.0881 | 118.4913 | 134.8437 | 114.4033 | 138.9318 | 3.227% | 6.455% | 9.682% | 120.3341 | 133.0009
Re, ppm | 0.0020185 | 0.00073947 | 0.00053957 | 0.0034975 | 0 | 0.0042369 | 36.63% | 73.27% | 109.9% | 0.0019176 | 0.0021194
S, wt.% | 4.2985 | 0.15868 | 3.9812 | 4.6159 | 3.8225 | 4.7746 | 3.691% | 7.383% | 11.07% | 4.0836 | 4.5135
Sb, ppm | 4.2156 | 0.24947 | 3.7167 | 4.7146 | 3.4672 | 4.964 | 5.918% | 11.84% | 17.75% | 4.0048 | 4.4264
Sc, ppm | 10.0955 | 0.57365 | 8.9482 | 11.2428 | 8.3745 | 11.8164 | 5.682% | 11.36% | 17.05% | 9.5907 | 10.6002
Se, ppm | 45.9473 | 3.4685 | 39.0103 | 52.8843 | 35.5418 | 56.3527 | 7.549% | 15.1% | 22.65% | 43.6499 | 48.2446
Sm, ppm | 5.7732 | 0.48508 | 4.803 | 6.7433 | 4.3179 | 7.2284 | 8.402% | 16.8% | 25.21% | 5.4845 | 6.0618
Sn, ppm | 37.4068 | 1.1323 | 35.1421 | 39.6715 | 34.0098 | 40.8039 | 3.027% | 6.054% | 9.081% | 35.5365 | 39.2772
Sr, ppm | 46.6102 | 2.4754 | 41.6595 | 51.5609 | 39.1841 | 54.0363 | 5.311% | 10.62% | 15.93% | 44.2797 | 48.9407
Ta, ppm | 0.19686 | 0.050858 | 0.095142 | 0.29858 | 0.044284 | 0.34943 | 25.83% | 51.67% | 77.5% | 0.18702 | 0.2067
Tb, ppm | 0.58842 | 0.040742 | 0.50694 | 0.6699 | 0.46619 | 0.71065 | 6.924% | 13.85% | 20.77% | 0.559 | 0.61784
Te, ppm | 0.171 | 0.023401 | 0.1242 | 0.21781 | 0.1008 | 0.24121 | 13.68% | 27.37% | 41.05% | 0.16245 | 0.17955
Th, ppm | 11.2869 | 0.94624 | 9.3944 | 13.1794 | 8.4482 | 14.1256 | 8.384% | 16.77% | 25.15% | 10.7226 | 11.8512
Ti, wt.% | 0.15031 | 0.014592 | 0.12113 | 0.1795 | 0.10654 | 0.19409 | 9.707% | 19.41% | 29.12% | 0.1428 | 0.15783
Tl, ppm | 0.70977 | 0.035748 | 0.63828 | 0.78127 | 0.60253 | 0.81701 | 5.037% | 10.07% | 15.11% | 0.67428 | 0.74526
Tm, ppm | 0.19633 | 0.0076292 | 0.18107 | 0.21159 | 0.17344 | 0.21921 | 3.886% | 7.772% | 11.66% | 0.18651 | 0.20614
U, ppm | 3.4555 | 0.19497 | 3.0655 | 3.8454 | 2.8706 | 4.0404 | 5.642% | 11.28% | 16.93% | 3.2827 | 3.6282
V, ppm | 138.195 | 6.3603 | 125.4743 | 150.9157 | 119.114 | 157.276 | 4.602% | 9.205% | 13.81% | 131.2853 | 145.1048
W, ppm | 18.7815 | 0.85148 | 17.0785 | 20.4845 | 16.2271 | 21.3359 | 4.534% | 9.067% | 13.6% | 17.8424 | 19.7206
Y, ppm | 13.5386 | 1.9099 | 9.7189 | 17.3584 | 7.809 | 19.2683 | 14.11% | 28.21% | 42.32% | 12.8617 | 14.2156
Yb, ppm | 1.3773 | 0.1466 | 1.0841 | 1.6705 | 0.9375 | 1.8171 | 10.64% | 21.29% | 31.93% | 1.3084 | 1.4462
Zn, ppm | 286.9808 | 10.8474 | 265.2861 | 308.6755 | 254.4387 | 319.5229 | 3.78% | 7.56% | 11.34% | 272.6318 | 301.3298
Zr, ppm | 62.1344 | 4.579 | 52.9765 | 71.2924 | 48.3975 | 75.8714 | 7.369% | 14.74% | 22.11% | 59.0277 | 65.2412
Aqua Regia Digestion
Ag, ppm | 17.1053 | 1.3263 | 14.4528 | 19.7578 | 13.1265 | 21.084 | 7.753% | 15.51% | 23.26% | 16.25 | 17.9605
Al, wt.% | 0.72738 | 0.059392 | 0.6086 | 0.84616 | 0.54921 | 0.90556 | 8.165% | 16.33% | 24.5% | 0.69101 | 0.76375
As, ppm | 97.1731 | 5.9779 | 85.2173 | 109.1289 | 79.2394 | 115.1068 | 6.152% | 12.3% | 18.46% | 92.3144 | 102.0317
Au, ppm | 0.017744 | 0.0030602 | 0.011623 | 0.023864 | 0.0085632 | 0.026924 | 17.25% | 34.49% | 51.74% | 0.016857 | 0.018631
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.43556 | 0.084984 | 0.26559 | 0.60552 | 0.1806 | 0.69051 | 19.51% | 39.02% | 58.53% | 0.41378 | 0.45733
Bi, ppm | 209.9302 | 15.6973 | 178.5357 | 241.3247 | 162.8384 | 257.022 | 7.477% | 14.95% | 22.43% | 199.4337 | 220.4267
Ca, wt.% | 0.078702 | 0.0030236 | 0.072655 | 0.084749 | 0.069631 | 0.087773 | 3.842% | 7.684% | 11.53% | 0.074767 | 0.082637
Cd, ppm | 0.61705 | 0.087249 | 0.44255 | 0.79155 | 0.3553 | 0.8788 | 14.14% | 28.28% | 42.42% | 0.5862 | 0.6479
Ce, ppm | 49.9306 | 2.6158 | 44.699 | 55.1623 | 42.0832 | 57.7781 | 5.239% | 10.48% | 15.72% | 47.4341 | 52.4272
Co, ppm | 29.5683 | 2.0676 | 25.4332 | 33.7035 | 23.3656 | 35.771 | 6.993% | 13.99% | 20.98% | 28.0899 | 31.0467
Cr, ppm | 21.0864 | 0.69764 | 19.6911 | 22.4816 | 18.9935 | 23.1793 | 3.308% | 6.617% | 9.925% | 20.032 | 22.1407
Cs, ppm | 2.0636 | 0.19319 | 1.6772 | 2.4499 | 1.484 | 2.6431 | 9.362% | 18.72% | 28.09% | 1.9604 | 2.1667
Cu, wt.% | 3.9481 | 0.12095 | 3.7062 | 4.19 | 3.5852 | 4.3109 | 3.064% | 6.127% | 9.191% | 3.7507 | 4.1455
Fe, wt.% | 8.2904 | 0.30864 | 7.6731 | 8.9077 | 7.3645 | 9.2163 | 3.723% | 7.446% | 11.17% | 7.8759 | 8.7049
Ga, ppm | 3.2486 | 0.39703 | 2.4545 | 4.0426 | 2.0575 | 4.4397 | 12.22% | 24.44% | 36.66% | 3.0862 | 3.411
Ge, ppm | 0.175 | 0.029652 | 0.1157 | 0.2343 | 0.086043 | 0.26396 | 16.94% | 33.89% | 50.83% | 0.16625 | 0.18375
Hf, ppm | 0.33978 | 0.021611 | 0.29655 | 0.383 | 0.27494 | 0.40461 | 6.36% | 12.72% | 19.08% | 0.32279 | 0.35676
Hg, ppm | 0.11402 | 0.025772 | 0.062478 | 0.16557 | 0.036706 | 0.19134 | 22.6% | 45.21% | 67.81% | 0.10832 | 0.11972
In, ppm | 3.2398 | 0.12406 | 2.9917 | 3.4879 | 2.8676 | 3.6119 | 3.829% | 7.659% | 11.49% | 3.0778 | 3.4018
K, wt.% | 0.19021 | 0.020477 | 0.14925 | 0.23116 | 0.12878 | 0.25164 | 10.77% | 21.53% | 32.3% | 0.1807 | 0.19972
La, ppm | 19.5236 | 2.0538 | 15.4159 | 23.6313 | 13.3621 | 25.6852 | 10.52% | 21.04% | 31.56% | 18.5475 | 20.4998
Li, ppm | 2.031 | 0.35388 | 1.3232 | 2.7388 | 0.96936 | 3.0926 | 17.42% | 34.85% | 52.27% | 1.9294 | 2.1325
Lu, ppm | 0.0915 | 0.01156 | 0.06838 | 0.11462 | 0.056819 | 0.12618 | 12.63% | 25.27% | 37.9% | 0.086925 | 0.096075
Mg, wt.% | 0.12776 | 0.012049 | 0.10366 | 0.15186 | 0.091616 | 0.16391 | 9.431% | 18.86% | 28.29% | 0.12137 | 0.13415
Mn, wt.% | 0.013706 | 0.00082151 | 0.012063 | 0.015349 | 0.011241 | 0.01617 | 5.994% | 11.99% | 17.98% | 0.013021 | 0.014391
Mo, ppm | 3.4024 | 0.23638 | 2.9296 | 3.8752 | 2.6933 | 4.1116 | 6.948% | 13.9% | 20.84% | 3.2323 | 3.5725
Na, wt.% | 0.010141 | 0.0016947 | 0.0067516 | 0.01353 | 0.0050568 | 0.015225 | 16.71% | 33.42% | 50.13% | 0.009634 | 0.010648
Nb, ppm | 0.055556 | 0.0087099 | 0.038136 | 0.072975 | 0.029426 | 0.081685 | 15.68% | 31.36% | 47.03% | 0.052778 | 0.058333
Ni, ppm | 10.2019 | 0.76752 | 8.6668 | 11.7369 | 7.8993 | 12.5044 | 7.523% | 15.05% | 22.57% | 9.6918 | 10.712
P, wt.% | 0.021728 | 0.0025058 | 0.016716 | 0.026739 | 0.01421 | 0.029245 | 11.53% | 23.07% | 34.6% | 0.020641 | 0.022814
Pb, ppm | 115.1026 | 6.0266 | 103.0494 | 127.1558 | 97.0228 | 133.1824 | 5.236% | 10.47% | 15.71% | 109.3475 | 120.8578
Rb, ppm | 13.7467 | 1.0616 | 11.6235 | 15.8699 | 10.562 | 16.9314 | 7.722% | 15.44% | 23.17% | 13.0594 | 14.434
Re, ppm | 0.0012813 | 0.00045524 | 0.00037076 | 0.0021917 | 0 | 0.002647 | 35.53% | 71.06% | 106.6% | 0.0012172 | 0.0013453
S, wt.% | 4.2498 | 0.21614 | 3.8175 | 4.6821 | 3.6013 | 4.8982 | 5.086% | 10.17% | 15.26% | 4.0373 | 4.4623
Sb, ppm | 2.9442 | 0.56051 | 1.8232 | 4.0653 | 1.2627 | 4.6258 | 19.04% | 38.07% | 57.11% | 2.797 | 3.0915
Sc, ppm | 1.4727 | 0.17308 | 1.1266 | 1.8189 | 0.95351 | 1.992 | 11.75% | 23.5% | 35.26% | 1.3991 | 1.5464
Se, ppm | 46.3087 | 3.0707 | 40.1672 | 52.4501 | 37.0964 | 55.5209 | 6.631% | 13.26% | 19.89% | 43.9932 | 48.6241
Sn, ppm | 34.5782 | 3.7243 | 27.1296 | 42.0269 | 23.4052 | 45.7512 | 10.77% | 21.54% | 32.31% | 32.8493 | 36.3071
Sr, ppm | 13.4925 | 2.4591 | 8.5743 | 18.4108 | 6.1152 | 20.8699 | 18.23% | 36.45% | 54.68% | 12.8179 | 14.1672
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.35897 | 0.048145 | 0.26268 | 0.45526 | 0.21453 | 0.5034 | 13.41% | 26.82% | 40.24% | 0.34102 | 0.37692
Te, ppm | 0.16403 | 0.013614 | 0.13681 | 0.19126 | 0.12319 | 0.20488 | 8.3% | 16.6% | 24.9% | 0.15583 | 0.17224
Th, ppm | 7.6785 | 0.57805 | 6.5224 | 8.8346 | 5.9444 | 9.4127 | 7.528% | 15.06% | 22.58% | 7.2946 | 8.0625
Ti, wt.% | 0.006777 | 0.0017909 | 0.0031952 | 0.010359 | 0.0014043 | 0.01215 | 26.43% | 52.85% | 79.28% | 0.0064382 | 0.0071159
Tl, ppm | 0.11103 | 0.010154 | 0.090722 | 0.13134 | 0.080568 | 0.14149 | 9.145% | 18.29% | 27.44% | 0.10548 | 0.11658
U, ppm | 2.0816 | 0.11446 | 1.8527 | 2.3106 | 1.7382 | 2.425 | 5.499% | 11% | 16.5% | 1.9776 | 2.1857
V, ppm | 18.7805 | 2.2919 | 14.1968 | 23.3642 | 11.9049 | 25.656 | 12.2% | 24.41% | 36.61% | 17.8415 | 19.7195
W, ppm | 13.6653 | 1.3052 | 11.0548 | 16.2758 | 9.7496 | 17.581 | 9.552% | 19.1% | 28.65% | 12.982 | 14.3486
Y, ppm | 5.5287 | 0.47132 | 4.5861 | 6.4714 | 4.1148 | 6.9427 | 8.525% | 17.05% | 25.57% | 5.2523 | 5.8052
Yb, ppm | 0.60808 | 0.026821 | 0.55444 | 0.66173 | 0.52762 | 0.68855 | 4.411% | 8.822% | 13.23% | 0.57768 | 0.63849
Zn, ppm | 271.94 | 7.2797 | 257.3805 | 286.4994 | 250.1008 | 293.7792 | 2.677% | 5.354% | 8.031% | 258.343 | 285.537
Zr, ppm | 10.7642 | 1.2505 | 8.2632 | 13.2653 | 7.0126 | 14.5158 | 11.62% | 23.23% | 34.85% | 10.226 | 11.3024
Borate Fusion XRF
Al2O3, wt.% | 10.8099 | 0.13737 | 10.5351 | 11.0846 | 10.3978 | 11.222 | 1.271% | 2.542% | 3.812% | 10.2694 | 11.3504
BaO, ppm | 2008.6475 | 144.1241 | 1720.3992 | 2296.8958 | 1576.2751 | 2441.0199 | 7.175% | 14.35% | 21.53% | 1908.2151 | 2109.0799
CaO, wt.% | 0.11503 | 0.0106 | 0.093827 | 0.13623 | 0.083227 | 0.14683 | 9.215% | 18.43% | 27.65% | 0.10928 | 0.12078
Cu, wt.% | 3.9506 | 0.10886 | 3.7329 | 4.1683 | 3.624 | 4.2772 | 2.756% | 5.511% | 8.267% | 3.7531 | 4.1481
Fe2O3, wt.% | 12.4208 | 0.19092 | 12.0389 | 12.8026 | 11.848 | 12.9935 | 1.537% | 3.074% | 4.611% | 11.7997 | 13.0418
K2O, wt.% | 2.5527 | 0.10235 | 2.348 | 2.7575 | 2.2457 | 2.8598 | 4.01% | 8.019% | 12.03% | 2.4251 | 2.6804
MgO, wt.% | 0.63185 | 0.067917 | 0.49602 | 0.76769 | 0.4281 | 0.83561 | 10.75% | 21.5% | 32.25% | 0.60026 | 0.66345
MnO, wt.% | 0.02226 | 0.0054808 | 0.011298 | 0.033222 | 0.0058177 | 0.038703 | 24.62% | 49.24% | 73.86% | 0.021147 | 0.023373
P2O5, wt.% | 0.069862 | 0.01359 | 0.042681 | 0.097043 | 0.029091 | 0.11063 | 19.45% | 38.91% | 58.36% | 0.066369 | 0.073355
S, wt.% | 4.3509 | 0.23353 | 3.8839 | 4.818 | 3.6503 | 5.0515 | 5.367% | 10.73% | 16.1% | 4.1334 | 4.5685
SiO2, wt.% | 62.3625 | 0.7077 | 60.9471 | 63.7779 | 60.2394 | 64.4856 | 1.135% | 2.27% | 3.404% | 59.2444 | 65.4806
TiO2, wt.% | 0.44094 | 0.02253 | 0.39588 | 0.486 | 0.37335 | 0.50853 | 5.11% | 10.22% | 15.33% | 0.41889 | 0.46299
Zn, ppm | 289.7671 | 24.9515 | 239.8641 | 339.67 | 214.9127 | 364.6215 | 8.611% | 17.22% | 25.83% | 275.2787 | 304.2554
Borate / Peroxide Fusion ICP
Al, wt.% | 5.7319 | 0.15968 | 5.4125 | 6.0512 | 5.2528 | 6.2109 | 2.786% | 5.572% | 8.358% | 5.4453 | 6.0185
As, ppm | 96.9533 | 9.5035 | 77.9463 | 115.9602 | 68.4428 | 125.4637 | 9.802% | 19.6% | 29.41% | 92.1056 | 101.8009
B, ppm | 113.7586 | 10.1783 | 93.4019 | 134.1152 | 83.2236 | 144.2935 | 8.947% | 17.89% | 26.84% | 108.0706 | 119.4465
Ba, ppm | 1739.3893 | 53.8913 | 1631.6067 | 1847.172 | 1577.7153 | 1901.0633 | 3.098% | 6.197% | 9.295% | 1652.4198 | 1826.3588
Be, ppm | 2.0569 | 0.1992 | 1.6585 | 2.4553 | 1.4593 | 2.6545 | 9.684% | 19.37% | 29.05% | 1.9541 | 2.1598
Bi, ppm | 215.9555 | 11.4726 | 193.0103 | 238.9007 | 181.5377 | 250.3733 | 5.312% | 10.62% | 15.94% | 205.1578 | 226.7533
Ca, wt.% | 0.096169 | 0.024604 | 0.046961 | 0.14538 | 0.022357 | 0.16998 | 25.58% | 51.17% | 76.75% | 0.09136 | 0.10098
Ce, ppm | 76.2399 | 4.0761 | 68.0878 | 84.392 | 64.0117 | 88.4681 | 5.346% | 10.69% | 16.04% | 72.4279 | 80.0519
Co, ppm | 28.3806 | 1.7228 | 24.935 | 31.8262 | 23.2122 | 33.549 | 6.07% | 12.14% | 18.21% | 26.9616 | 29.7996
Cr, ppm | 91.1811 | 18.3909 | 54.3994 | 127.9629 | 36.0085 | 146.3538 | 20.17% | 40.34% | 60.51% | 86.6221 | 95.7402
Cs, ppm | 4.9188 | 0.24788 | 4.423 | 5.4145 | 4.1751 | 5.6624 | 5.039% | 10.08% | 15.12% | 4.6728 | 5.1647
Cu, wt.% | 3.9577 | 0.091512 | 3.7747 | 4.1408 | 3.6832 | 4.2323 | 2.312% | 4.624% | 6.937% | 3.7598 | 4.1556
Dy, ppm | 4.1169 | 0.19255 | 3.7318 | 4.5021 | 3.5393 | 4.6946 | 4.677% | 9.354% | 14.03% | 3.9111 | 4.3228
Er, ppm | 2.362 | 0.15154 | 2.0589 | 2.6651 | 1.9073 | 2.8166 | 6.416% | 12.83% | 19.25% | 2.2439 | 2.4801
Eu, ppm | 1.2561 | 0.12028 | 1.0156 | 1.4967 | 0.89529 | 1.617 | 9.575% | 19.15% | 28.73% | 1.1933 | 1.3189
Fe, wt.% | 8.6761 | 0.13847 | 8.3991 | 8.953 | 8.2606 | 9.0915 | 1.596% | 3.192% | 4.788% | 8.2423 | 9.1099
Ga, ppm | 18.3076 | 2.0598 | 14.1879 | 22.4273 | 12.128 | 24.4871 | 11.25% | 22.5% | 33.75% | 17.3922 | 19.223
Gd, ppm | 4.9416 | 0.3326 | 4.2764 | 5.6068 | 3.9438 | 5.9394 | 6.731% | 13.46% | 20.19% | 4.6945 | 5.1887
Ge, ppm | 2.3277 | 0.48297 | 1.3618 | 3.2936 | 0.87878 | 3.7766 | 20.75% | 41.5% | 62.25% | 2.2113 | 2.4441
Ho, ppm | 0.80914 | 0.054091 | 0.70096 | 0.91732 | 0.64687 | 0.97141 | 6.685% | 13.37% | 20.05% | 0.76868 | 0.8496
In, ppm | 3.3204 | 0.19211 | 2.9361 | 3.7046 | 2.744 | 3.8967 | 5.786% | 11.57% | 17.36% | 3.1543 | 3.4864
K, wt.% | 2.2426 | 0.071872 | 2.0988 | 2.3863 | 2.027 | 2.4582 | 3.205% | 6.41% | 9.615% | 2.1304 | 2.3547
La, ppm | 37.1848 | 2.8358 | 31.5132 | 42.8563 | 28.6774 | 45.6921 | 7.626% | 15.25% | 22.88% | 35.3255 | 39.044
Li, ppm | 17.7891 | 1.9687 | 13.8517 | 21.7265 | 11.883 | 23.6952 | 11.07% | 22.13% | 33.2% | 16.8996 | 18.6785
Lu, ppm | 0.338 | 0.022697 | 0.29261 | 0.3834 | 0.26991 | 0.40609 | 6.715% | 13.43% | 20.15% | 0.3211 | 0.3549
Mg, wt.% | 0.37557 | 0.015245 | 0.34508 | 0.40606 | 0.32983 | 0.42131 | 4.059% | 8.119% | 12.18% | 0.35679 | 0.39435
Mn, wt.% | 0.017503 | 0.0019669 | 0.013569 | 0.021437 | 0.011602 | 0.023404 | 11.24% | 22.47% | 33.71% | 0.016628 | 0.018378
Mo, ppm | 3.7415 | 0.80545 | 2.1306 | 5.3524 | 1.3252 | 6.1579 | 21.53% | 43.05% | 64.58% | 3.5545 | 3.9286
Nb, ppm | 8.3708 | 1.364 | 5.6429 | 11.0988 | 4.2789 | 12.4627 | 16.29% | 32.59% | 48.88% | 7.9523 | 8.7894
Nd, ppm | 31.6456 | 1.0721 | 29.5013 | 33.7899 | 28.4292 | 34.862 | 3.388% | 6.776% | 10.16% | 30.0633 | 33.2279
P, wt.% | 0.029566 | 0.004104 | 0.021358 | 0.037774 | 0.017254 | 0.041878 | 13.88% | 27.76% | 41.64% | 0.028087 | 0.031044
Pb, ppm | 120.954 | 20.3554 | 80.2433 | 161.6647 | 59.8879 | 182.0201 | 16.83% | 33.66% | 50.49% | 114.9063 | 127.0017
Pr, ppm | 8.4529 | 0.3597 | 7.7335 | 9.1723 | 7.3738 | 9.532 | 4.255% | 8.511% | 12.77% | 8.0302 | 8.8755
Rb, ppm | 123.8695 | 5.983 | 111.9035 | 135.8354 | 105.9205 | 141.8184 | 4.83% | 9.66% | 14.49% | 117.676 | 130.063
Re, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 4.4236 | 0.1712 | 4.0812 | 4.766 | 3.91 | 4.9372 | 3.87% | 7.74% | 11.61% | 4.2025 | 4.6448
Sb, ppm | 3.9898 | 0.72552 | 2.5388 | 5.4409 | 1.8133 | 6.1664 | 18.18% | 36.37% | 54.55% | 3.7903 | 4.1893
Sc, ppm | 9.4436 | 1.2276 | 6.9885 | 11.8987 | 5.7609 | 13.1262 | 13% | 26% | 39% | 8.9714 | 9.9158
Se, ppm | 46.2437 | 4.219 | 37.8056 | 54.6818 | 33.5866 | 58.9008 | 9.123% | 18.25% | 27.37% | 43.9315 | 48.5559
Si, wt.% | 29.7281 | 0.8511 | 28.0259 | 31.4303 | 27.1748 | 32.2814 | 2.863% | 5.726% | 8.589% | 28.2417 | 31.2145
Sm, ppm | 5.8511 | 0.27826 | 5.2946 | 6.4076 | 5.0163 | 6.6859 | 4.756% | 9.512% | 14.27% | 5.5585 | 6.1436
Sn, ppm | 45.0636 | 5.0124 | 35.0387 | 55.0884 | 30.0263 | 60.1009 | 11.12% | 22.25% | 33.37% | 42.8104 | 47.3168
Sr, ppm | 53.9512 | 5.1913 | 43.5686 | 64.3338 | 38.3773 | 69.5251 | 9.622% | 19.24% | 28.87% | 51.2537 | 56.6488
Ta, ppm | 0.94082 | 0.34999 | 0.24085 | 1.6408 | 0 | 1.9908 | 37.2% | 74.4% | 111.6% | 0.89378 | 0.98786
Tb, ppm | 0.71784 | 0.040337 | 0.63717 | 0.79851 | 0.59683 | 0.83885 | 5.619% | 11.24% | 16.86% | 0.68195 | 0.75373
Th, ppm | 11.69 | 1.2492 | 9.1915 | 14.1885 | 7.9423 | 15.4378 | 10.69% | 21.37% | 32.06% | 11.1055 | 12.2745
Ti, wt.% | 0.26527 | 0.0075899 | 0.25009 | 0.28045 | 0.2425 | 0.28804 | 2.861% | 5.722% | 8.584% | 0.252 | 0.27853
Tl, ppm | 0.74349 | 0.051551 | 0.64039 | 0.84659 | 0.58883 | 0.89814 | 6.934% | 13.87% | 20.8% | 0.70631 | 0.78066
Tm, ppm | 0.3442 | 0.029789 | 0.28462 | 0.40378 | 0.25483 | 0.43357 | 8.655% | 17.31% | 25.96% | 0.32699 | 0.36141
U, ppm | 3.8007 | 0.32147 | 3.1577 | 4.4436 | 2.8363 | 4.7651 | 8.458% | 16.92% | 25.37% | 3.6106 | 3.9907
V, ppm | 144.0234 | 5.8173 | 132.3887 | 155.6581 | 126.5714 | 161.4754 | 4.039% | 8.078% | 12.12% | 136.8222 | 151.2246
W, ppm | 17.8191 | 2.6341 | 12.5509 | 23.0873 | 9.9168 | 25.7214 | 14.78% | 29.56% | 44.35% | 16.9282 | 18.7101
Y, ppm | 23.5837 | 0.91513 | 21.7534 | 25.4139 | 20.8383 | 26.3291 | 3.88% | 7.761% | 11.64% | 22.4045 | 24.7629
Yb, ppm | 2.2463 | 0.1839 | 1.8785 | 2.6141 | 1.6946 | 2.798 | 8.187% | 16.37% | 24.56% | 2.134 | 2.3586
Zn, ppm | 283.3654 | 22.7162 | 237.933 | 328.7979 | 215.2167 | 351.5141 | 8.017% | 16.03% | 24.05% | 269.1972 | 297.5337
Infrared Combustion
S, wt.% | 4.4299 | 0.11264 | 4.2046 | 4.6552 | 4.092 | 4.7678 | 2.543% | 5.085% | 7.628% | 4.2084 | 4.6514
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Uncertainty & Tolerance Limits
Table 2. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 931b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
4-Acid Digestion
Ag, Silver (ppm) | 16.8756 | 16.2657 | 17.4854 | 16.5234 | 17.2278
Al, Aluminium (wt.%) | 5.5432 | 5.4168 | 5.6697 | 5.4222 | 5.6642
As, Arsenic (ppm) | 98.5527 | 94.1955 | 102.9099 | 96.1227 | 100.9827
Be, Beryllium (ppm) | 2.0818 | 1.9391 | 2.2244 | 2.0274 | 2.1361
Bi, Bismuth (ppm) | 211.3386 | 200.7189 | 221.9583 | 205.443 | 217.2341
Ca, Calcium (wt.%) | 0.083318 | 0.07912 | 0.087515 | 0.080757 | 0.085878
Cd, Cadmium (ppm) | 0.64385 | 0.58495 | 0.70275 | 0.60861 | 0.67909
Ce, Cerium (ppm) | 70.6165 | 65.7752 | 75.4578 | 67.577 | 73.656
Co, Cobalt (ppm) | 29.5646 | 28.2644 | 30.8649 | 28.8841 | 30.2451
Cr, Chromium (ppm) | 72.1313 | 67.1163 | 77.1462 | 69.404 | 74.8586
Cs, Caesium (ppm) | 4.7453 | 4.5219 | 4.9687 | 4.6088 | 4.8818
Cu, Copper (wt.%) | 3.9471 | 3.8388 | 4.0554 | 3.8776 | 4.0167
Dy, Dysprosium (ppm) | 3.0608 | 2.7433 | 3.3783 | 2.9113 | 3.2103
Er, Erbium (ppm) | 1.3743 | 1.2636 | 1.4849 | 1.2989 | 1.4496
Eu, Europium (ppm) | 1.0906 | 0.95929 | 1.222 | 1.0363 | 1.145
Fe, Iron (wt.%) | 8.552 | 8.2739 | 8.8301 | 8.399 | 8.7049
Ga, Gallium (ppm) | 17.7095 | 16.7219 | 18.6971 | 17.1165 | 18.3026
Gd, Gadolinium (ppm) | 4.4278 | 3.8591 | 4.9964 | 4.2021 | 4.6535
Ge, Germanium (ppm) | 0.20952 | 0.1675 | 0.25155 | 0.18492 | 0.23412
Hf, Hafnium (ppm) | 1.8155 | 1.7267 | 1.9044 | 1.7542 | 1.8769
Ho, Holmium (ppm) | 0.48863 | 0.44387 | 0.53339 | 0.46376 | 0.5135
In, Indium (ppm) | 3.3152 | 3.1662 | 3.4643 | 3.2128 | 3.4177
K, Potassium (wt.%) | 2.1761 | 2.1043 | 2.248 | 2.1338 | 2.2185
La, Lanthanum (ppm) | 32.3429 | 29.9658 | 34.72 | 31.0324 | 33.6534
Li, Lithium (ppm) | 16.9919 | 16.1067 | 17.877 | 16.5147 | 17.469
Lu, Lutetium (ppm) | 0.22246 | 0.19462 | 0.2503 | 0.20573 | 0.2392
Mg, Magnesium (wt.%) | 0.35396 | 0.33902 | 0.3689 | 0.34404 | 0.36388
Mn, Manganese (wt.%) | 0.016092 | 0.015601 | 0.016583 | 0.015675 | 0.016509
Mo, Molybdenum (ppm) | 3.3832 | 3.2165 | 3.5498 | 3.2352 | 3.5311
Na, Sodium (wt.%) | 0.076484 | 0.073823 | 0.079146 | 0.073297 | 0.079671
Nb, Niobium (ppm) | 2.7246 | 2.3713 | 3.0778 | 2.5508 | 2.8983
Nd, Neodymium (ppm) | 29.8611 | 27.0323 | 32.6899 | 28.9652 | 30.7571
Ni, Nickel (ppm) | 12.5914 | 11.8966 | 13.2862 | 12.0701 | 13.1128
P, Phosphorus (wt.%) | 0.027157 | 0.025873 | 0.028441 | 0.02633 | 0.027984
Pb, Lead (ppm) | 121.0905 | 114.8673 | 127.3138 | 117.6638 | 124.5173
Pr, Praseodymium (ppm) | 8.1238 | 7.2756 | 8.9719 | 7.8949 | 8.3527
Rb, Rubidium (ppm) | 126.6675 | 122.2415 | 131.0935 | 123.9252 | 129.4099
Re, Rhenium (ppm) | 0.0020185 | 0.00073903 | 0.003298 | IND | IND
S, Sulphur (wt.%) | 4.2985 | 4.1951 | 4.402 | 4.2186 | 4.3785
Sb, Antimony (ppm) | 4.2156 | 3.9743 | 4.457 | 4.0405 | 4.3908
Sc, Scandium (ppm) | 10.0955 | 9.6063 | 10.5846 | 9.8075 | 10.3834
Se, Selenium (ppm) | 45.9473 | 43.1639 | 48.7307 | 43.939 | 47.9556
Sm, Samarium (ppm) | 5.7732 | 5.2756 | 6.2707 | 5.5004 | 6.0459
Sn, Tin (ppm) | 37.4068 | 35.6187 | 39.195 | 36.4887 | 38.3249
Sr, Strontium (ppm) | 46.6102 | 44.3315 | 48.8889 | 44.9557 | 48.2647
Ta, Tantalum (ppm) | 0.19686 | 0.15679 | 0.23693 | 0.15796 | 0.23576
Tb, Terbium (ppm) | 0.58842 | 0.52329 | 0.65355 | 0.55507 | 0.62177
Te, Tellurium (ppm) | 0.171 | 0.13665 | 0.20536 | IND | IND
Th, Thorium (ppm) | 11.2869 | 10.552 | 12.0218 | 10.9071 | 11.6667
Ti, Titanium (wt.%) | 0.15031 | 0.14144 | 0.15918 | 0.14566 | 0.15497
Tl, Thallium (ppm) | 0.70977 | 0.67203 | 0.74751 | 0.68629 | 0.73325
Tm, Thulium (ppm) | 0.19633 | 0.18597 | 0.20668 | IND | IND
U, Uranium (ppm) | 3.4555 | 3.2783 | 3.6326 | 3.3511 | 3.5598
V, Vanadium (ppm) | 138.195 | 132.9877 | 143.4024 | 135.2293 | 141.1607
W, Tungsten (ppm) | 18.7815 | 17.9704 | 19.5926 | 18.1787 | 19.3843
Y, Yttrium (ppm) | 13.5386 | 12.3853 | 14.692 | 12.9833 | 14.094
Yb, Ytterbium (ppm) | 1.3773 | 1.2272 | 1.5274 | 1.2871 | 1.4675
Zn, Zinc (ppm) | 286.9808 | 275.94 | 298.0216 | 281.5807 | 292.3809
Zr, Zirconium (ppm) | 62.1344 | 58.9026 | 65.3663 | 60.3356 | 63.9333
Aqua Regia Digestion
Ag, Silver (ppm) | 17.1053 | 16.3911 | 17.8195 | 16.5827 | 17.6279
Al, Aluminium (wt.%) | 0.72738 | 0.68872 | 0.76604 | 0.70769 | 0.74707
As, Arsenic (ppm) | 97.1731 | 93.3596 | 100.9866 | 94.5366 | 99.8095
Au, Gold (ppm) | 0.017744 | 0.014645 | 0.020842 | 0.015749 | 0.019738
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.43556 | 0.36762 | 0.5035 | 0.41411 | 0.457
Bi, Bismuth (ppm) | 209.9302 | 197.6789 | 222.1815 | 203.944 | 215.9164
Ca, Calcium (wt.%) | 0.078702 | 0.076494 | 0.08091 | 0.0764 | 0.081004
Cd, Cadmium (ppm) | 0.61705 | 0.55491 | 0.6792 | 0.58658 | 0.64752
Ce, Cerium (ppm) | 49.9306 | 47.7448 | 52.1165 | 48.8533 | 51.008
Co, Cobalt (ppm) | 29.5683 | 28.2949 | 30.8418 | 28.827 | 30.3096
Cr, Chromium (ppm) | 21.0864 | 19.6828 | 22.49 | 20.3468 | 21.826
Cs, Caesium (ppm) | 2.0636 | 1.9336 | 2.1935 | 2.0094 | 2.1177
Cu, Copper (wt.%) | 3.9481 | 3.8488 | 4.0473 | 3.8864 | 4.0097
Fe, Iron (wt.%) | 8.2904 | 8.0626 | 8.5182 | 8.1402 | 8.4405
Ga, Gallium (ppm) | 3.2486 | 2.9965 | 3.5007 | 3.1201 | 3.3771
Ge, Germanium (ppm) | 0.175 | 0.1405 | 0.2095 | IND | IND
Hf, Hafnium (ppm) | 0.33978 | 0.31365 | 0.3659 | 0.31992 | 0.35963
Hg, Mercury (ppm) | 0.11402 | 0.091987 | 0.13606 | 0.088681 | 0.13936
In, Indium (ppm) | 3.2398 | 3.1347 | 3.3448 | 3.1645 | 3.3151
K, Potassium (wt.%) | 0.19021 | 0.1769 | 0.20351 | 0.18215 | 0.19826
La, Lanthanum (ppm) | 19.5236 | 18.0698 | 20.9775 | 18.7201 | 20.3272
Li, Lithium (ppm) | 2.031 | 1.7434 | 2.3186 | 1.8596 | 2.2024
Lu, Lutetium (ppm) | 0.0915 | 0.071414 | 0.11159 | IND | IND
Mg, Magnesium (wt.%) | 0.12776 | 0.1183 | 0.13723 | 0.12356 | 0.13197
Mn, Manganese (wt.%) | 0.013706 | 0.013123 | 0.014288 | 0.013422 | 0.01399
Mo, Molybdenum (ppm) | 3.4024 | 3.2243 | 3.5805 | 3.2624 | 3.5424
Na, Sodium (wt.%) | 0.010141 | 0.0088629 | 0.011419 | IND | IND
Nb, Niobium (ppm) | 0.055556 | 0.043455 | 0.067657 | IND | IND
Ni, Nickel (ppm) | 10.2019 | 9.6264 | 10.7774 | 9.8064 | 10.5974
P, Phosphorus (wt.%) | 0.021728 | 0.020262 | 0.023194 | 0.021012 | 0.022444
Pb, Lead (ppm) | 115.1026 | 111.273 | 118.9322 | 112.0549 | 118.1504
Rb, Rubidium (ppm) | 13.7467 | 12.8886 | 14.6048 | 13.113 | 14.3804
Re, Rhenium (ppm) | 0.0012813 | 0.00031279 | 0.0022497 | IND | IND
S, Sulphur (wt.%) | 4.2498 | 4.1105 | 4.389 | 4.1533 | 4.3463
Sb, Antimony (ppm) | 2.9442 | 2.6416 | 3.2469 | 2.8238 | 3.0647
Sc, Scandium (ppm) | 1.4727 | 1.3362 | 1.6093 | IND | IND
Se, Selenium (ppm) | 46.3087 | 44.1546 | 48.4627 | 44.8181 | 47.7992
Sn, Tin (ppm) | 34.5782 | 32.3984 | 36.7581 | 33.4641 | 35.6923
Sr, Strontium (ppm) | 13.4925 | 12.0014 | 14.9837 | 13.0283 | 13.9568
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.35897 | 0.28864 | 0.4293 | 0.32807 | 0.38986
Te, Tellurium (ppm) | 0.16403 | 0.14238 | 0.18569 | 0.10217 | 0.2259
Th, Thorium (ppm) | 7.6785 | 7.1515 | 8.2056 | 7.4381 | 7.919
Ti, Titanium (wt.%) | 0.006777 | 0.0056479 | 0.0079061 | 0.0061703 | 0.0073837
Tl, Thallium (ppm) | 0.11103 | 0.098708 | 0.12335 | IND | IND
U, Uranium (ppm) | 2.0816 | 1.9774 | 2.1859 | 2.0171 | 2.1462
V, Vanadium (ppm) | 18.7805 | 17.488 | 20.0729 | 17.9485 | 19.6124
W, Tungsten (ppm) | 13.6653 | 12.7089 | 14.6217 | 13.0964 | 14.2342
Y, Yttrium (ppm) | 5.5287 | 5.1999 | 5.8576 | 5.3536 | 5.7039
Yb, Ytterbium (ppm) | 0.60808 | 0.57801 | 0.63815 | IND | IND
Zn, Zinc (ppm) | 271.94 | 265.1411 | 278.7389 | 267.0326 | 276.8473
Zr, Zirconium (ppm) | 10.7642 | 9.9834 | 11.5451 | 10.4121 | 11.1163
Borate Fusion XRF
Al2O3, Aluminium(III) oxide (wt.%) | 10.8099 | 10.6681 | 10.9516 | 10.6776 | 10.9421
BaO, Barium oxide (ppm) | 2008.6475 | 1877.2501 | 2140.0449 | 1879.2746 | 2138.0204
CaO, Calcium oxide (wt.%) | 0.11503 | 0.10434 | 0.12571 | IND | IND
Cu, Copper (wt.%) | 3.9506 | 3.8597 | 4.0414 | 3.9028 | 3.9984
Fe2O3, Iron(III) oxide (wt.%) | 12.4208 | 12.2358 | 12.6058 | 12.2913 | 12.5502
K2O, Potassium oxide (wt.%) | 2.5527 | 2.4098 | 2.6957 | 2.5182 | 2.5873
MgO, Magnesium oxide (wt.%) | 0.63185 | 0.53415 | 0.72956 | 0.59032 | 0.67339
MnO, Manganese oxide (wt.%) | 0.02226 | 0.015433 | 0.029088 | IND | IND
P2O5, Phosphorus(V) oxide (wt.%) | 0.069862 | 0.049435 | 0.090289 | IND | IND
S, Sulphur (wt.%) | 4.3509 | 4.1769 | 4.525 | 4.2844 | 4.4174
SiO2, Silicon dioxide (wt.%) | 62.3625 | 61.6568 | 63.0681 | 61.7745 | 62.9505
TiO2, Titanium dioxide (wt.%) | 0.44094 | 0.41061 | 0.47127 | 0.4263 | 0.45558
Zn, Zinc (ppm) | 289.7671 | 267.7307 | 311.8034 | 273.7577 | 305.7765
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.7319 | 5.5973 | 5.8665 | 5.6375 | 5.8262
As, Arsenic (ppm) | 96.9533 | 87.7007 | 106.2058 | 93.2591 | 100.6474
B, Boron (ppm) | 113.7586 | 97.1387 | 130.3784 | 100.6996 | 126.8175
Ba, Barium (ppm) | 1739.3893 | 1672.3707 | 1806.4079 | 1690.8429 | 1787.9357
Be, Beryllium (ppm) | 2.0569 | 1.9847 | 2.1291 | IND | IND
Bi, Bismuth (ppm) | 215.9555 | 203.4975 | 228.4136 | 208.6124 | 223.2987
Ca, Calcium (wt.%) | 0.096169 | 0.071073 | 0.12126 | IND | IND
Ce, Cerium (ppm) | 76.2399 | 71.009 | 81.4707 | 72.6082 | 79.8716
Co, Cobalt (ppm) | 28.3806 | 26.9456 | 29.8156 | 27.0747 | 29.6865
Cr, Chromium (ppm) | 91.1811 | 82.1705 | 100.1917 | 84.1191 | 98.2432
Cs, Caesium (ppm) | 4.9188 | 4.604 | 5.2335 | 4.6652 | 5.1723
Cu, Copper (wt.%) | 3.9577 | 3.8426 | 4.0728 | 3.8752 | 4.0403
Dy, Dysprosium (ppm) | 4.1169 | 3.8989 | 4.335 | 3.8799 | 4.354
Er, Erbium (ppm) | 2.362 | 2.0605 | 2.6634 | 2.2415 | 2.4824
Eu, Europium (ppm) | 1.2561 | 1.1018 | 1.4105 | 1.1449 | 1.3673
Fe, Iron (wt.%) | 8.6761 | 8.4879 | 8.8642 | 8.5491 | 8.803
Ga, Gallium (ppm) | 18.3076 | 16.7276 | 19.8876 | 17.2166 | 19.3986
Gd, Gadolinium (ppm) | 4.9416 | 4.4768 | 5.4064 | 4.6449 | 5.2383
Ge, Germanium (ppm) | 2.3277 | 1.6915 | 2.9639 | 1.95 | 2.7054
Ho, Holmium (ppm) | 0.80914 | 0.74084 | 0.87744 | 0.76062 | 0.85766
In, Indium (ppm) | 3.3204 | 3.0774 | 3.5634 | 3.1471 | 3.4936
K, Potassium (wt.%) | 2.2426 | 2.1785 | 2.3067 | 2.1856 | 2.2995
La, Lanthanum (ppm) | 37.1848 | 34.4373 | 39.9322 | 35.4409 | 38.9287
Li, Lithium (ppm) | 17.7891 | 14.6017 | 20.9764 | IND | IND
Lu, Lutetium (ppm) | 0.338 | 0.29983 | 0.37617 | 0.30607 | 0.36994
Mg, Magnesium (wt.%) | 0.37557 | 0.36039 | 0.39075 | 0.36262 | 0.38852
Mn, Manganese (wt.%) | 0.017503 | 0.016532 | 0.018473 | 0.017056 | 0.01795
Mo, Molybdenum (ppm) | 3.7415 | 2.8876 | 4.5955 | IND | IND
Nb, Niobium (ppm) | 8.3708 | 6.6984 | 10.0433 | 7.0016 | 9.7401
Nd, Neodymium (ppm) | 31.6456 | 29.7192 | 33.572 | 30.2779 | 33.0133
P, Phosphorus (wt.%) | 0.029566 | 0.023852 | 0.03528 | IND | IND
Pb, Lead (ppm) | 120.954 | 105.6212 | 136.2868 | 114.4651 | 127.4429
Pr, Praseodymium (ppm) | 8.4529 | 8.0146 | 8.8912 | 8.0426 | 8.8632
Rb, Rubidium (ppm) | 123.8695 | 117.8357 | 129.9033 | 118.5071 | 129.2319
Re, Rhenium (ppm) | < 0.1 | IND | IND | IND | IND
S, Sulphur (wt.%) | 4.4236 | 4.2804 | 4.5669 | 4.3171 | 4.5302
Sb, Antimony (ppm) | 3.9898 | 3.139 | 4.8406 | 3.7163 | 4.2634
Sc, Scandium (ppm) | 9.4436 | 7.9698 | 10.9174 | IND | IND
Se, Selenium (ppm) | 46.2437 | 41.6959 | 50.7915 | 42.3628 | 50.1247
Si, Silicon (wt.%) | 29.7281 | 29.1135 | 30.3427 | 29.0091 | 30.4471
Sm, Samarium (ppm) | 5.8511 | 5.5635 | 6.1387 | 5.5862 | 6.116
Sn, Tin (ppm) | 45.0636 | 39.8347 | 50.2925 | 41.5272 | 48.6
Sr, Strontium (ppm) | 53.9512 | 50.0067 | 57.8957 | 51.8701 | 56.0323
Ta, Tantalum (ppm) | 0.94082 | 0.68323 | 1.1984 | IND | IND
Tb, Terbium (ppm) | 0.71784 | 0.66796 | 0.76771 | 0.68284 | 0.75284
Th, Thorium (ppm) | 11.69 | 10.7171 | 12.663 | 10.982 | 12.3981
Ti, Titanium (wt.%) | 0.26527 | 0.25753 | 0.27301 | 0.25738 | 0.27315
Tl, Thallium (ppm) | 0.74349 | 0.65344 | 0.83353 | IND | IND
Tm, Thulium (ppm) | 0.3442 | 0.30309 | 0.38531 | 0.31659 | 0.37181
U, Uranium (ppm) | 3.8007 | 3.5103 | 4.0911 | 3.6453 | 3.956
V, Vanadium (ppm) | 144.0234 | 138.3335 | 149.7133 | 139.4985 | 148.5483
W, Tungsten (ppm) | 17.8191 | 15.5456 | 20.0927 | 16.4285 | 19.2098
Y, Yttrium (ppm) | 23.5837 | 22.587 | 24.5804 | 22.6106 | 24.5567
Yb, Ytterbium (ppm) | 2.2463 | 1.9983 | 2.4943 | 2.1485 | 2.3441
Zn, Zinc (ppm) | 283.3654 | 269.1372 | 297.5937 | 272.1555 | 294.5754
Infrared Combustion
S, Sulphur (wt.%) | 4.4299 | 4.3534 | 4.5065 | 4.3775 | 4.4823
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 3. Indicative Values for OREAS 931b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
4-Acid Digestion
B | ppm | 51.3333 | Ba | ppm | 581.1503 | Hg | ppm | 0.15515
Aqua Regia Digestion
Ba | ppm | 91.9789 | Gd | ppm | 2.8962 | Pr | ppm | 5.4337
Dy | ppm | 1.6771 | Ho | ppm | 0.28382 | Pt | ppb | < 5
Er | ppm | 0.66517 | Nd | ppm | 16.9967 | Sm | ppm | 3.8337
Eu | ppm | 0.60225 | Pd | ppb | 149.8333 | Tm | ppm | 0.090083
Borate Fusion XRF
Bi | ppm | 212.0613 | Mo | ppm | < 33 | Sn | ppm | < 160
Ce | ppm | < 80 | Na2O | wt.% | 0.15426 | SrO | ppm | 51.4444
Cl | ppm | < 331.815 | Nb | ppm | < 59.95051 | Ta | ppm | < 80
Co | ppm | < 60 | Nd | ppm | < 90 | Th | ppm | < 44
Cr2O3 | ppm | 128.9203 | Ni | ppm | 201.0917 | TOT_XRF | wt.% | 99.9167
F | ppm | < 3000 | Pb | ppm | 136.2155 | U | ppm | < 42
Ga | ppm | < 70 | Pr | ppm | < 80 | V2O5 | ppm | 281.8598
Gd | ppm | < 90 | Rb | ppm | 139.8011 | W | ppm | < 80
Hf | ppm | < 80 | Sb | ppm | < 170 | Y | ppm | < 39
La | ppm | < 170 | Sm | ppm | < 170 | ZrO2 | ppm | 45.0333
Borate / Peroxide Fusion ICP
Ag | ppm | 17.1778 | Na | wt.% | 0.080366 | Zr | ppm | 82.195
Cd | ppm | 0.63889 | Ni | ppm | 21.8668
Hf | ppm | 2.2945 | Te | ppm | < 1
Thermogravimetry
LOI1000 | wt.% | 5.2115
Infrared Combustion
C | wt.% | 0.67919
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 931b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*OES | lithium borate fusion with inductively coupled plasma optical emission spectroscopy
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
IRC | infrared combustion
OxBF*XRF | oxidising lithium borate fusion with X-ray fluorescence spectroscopy
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 931b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
MSALABS, Vancouver, BC, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS, Randfontein, Gauteng, South Africa
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
UIS Analytical Services, Centurion , South Africa

## 4-Acid
| Analytical results for Ag in OREAS 931b (Certified Value 16.9 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 16.9 | 17.763 | 16.61 | 14.6 | 16.3 | 16.15 | 16.6 | 16.85 | 16.75 | 15.88 | 14.6532 | > 10 | 14.77 | 17.2 | 16.71 | 16.7 | 17.1 | 17.64 | 16.6724 | 17.1 | 19.7 | 17 | 16.255 | 1
| 1 | 2 | 17 | 17.9857 | 16.64 | 14.9 | 17.25 | 16.3 | 16.6 | 16.8 | 17.15 | 16.76 | 15.4622 | > 10 | 14.39 | 17.2 | 17.4 | 16.7 | 16.7 | 17.17 | 16.7071 | 17.01 | 19.1 | 17 | 16.3519 | 19
| 1 | 3 | 17.2 | 17.9852 | 16.57 | 14.9 | 16.9 | 17.1 | 15.6 | 16 | 17.25 | 15.61 | 15.0533 | > 10 | 14.85 | 17 | 16.2 | 17 | 16.6 | 17.21 | 16.7745 | 16.42 | 19.6 | 17 | 16.8571 | 16
| 1 | 4 | 17.2 | 17.9866 | 16.6 | 14.7 | 16.3 | 17.1 | 15.8 | 17 | 17 | 16.06 | 15.742 | > 10 | 14.43 | 17.9 | 16.98 | 16.6 | 17.5 | 17.67 | 17.0818 | 16.88 | 19.8 | 17 | 16.7359 | 16.8756
| 1 | 5 | 16.9 | 18.0509 | 16.62 | 15.6 | 16.8 | 17.45 | 15.25 | 16.7 | 17.3 | 16.29 | 15.4655 | > 10 | 14.38 | 17.6 | 16.79 | 16.4 | 17.9 | 16.84 | 16.8162 | 17.38 | 19.9 | 19 | 16.3448 | 7
| 1 | 6 | 17 | 17.7835 | 16.5 | 14.9 | 16.75 | 17.3 | 15.5 | 16.8 | 17 | 15.88 | 15.6656 | > 10 | 14.25 | 17.3 | 16.78 | 16.7 | 17.6 | 16.9 | 17.1412 | 17.09 | 19.6 | 17 | 16.7007
| Mean |  | 17.0333 | 17.9258 | 16.59 | 14.9333 | 16.7167 | 16.9 | 15.8917 | 16.6917 | 17.075 | 16.08 | 15.3403 |  | 14.5117 | 17.3667 | 16.81 | 16.6833 | 17.2333 | 17.2383 | 16.8655 | 16.98 | 19.6167 | 17.3333 | 16.5409
| Median |  | 17 | 17.9854 | 16.605 | 14.9 | 16.775 | 17.1 | 15.7 | 16.8 | 17.075 | 15.97 | 15.4639 |  | 14.41 | 17.25 | 16.785 | 16.7 | 17.3 | 17.19 | 16.7954 | 17.05 | 19.65 | 17 | 16.5263
| Std Dev. |  | 0.13663 | 0.12099 | 0.0498 | 0.35024 | 0.36697 | 0.54129 | 0.57656 | 0.35273 | 0.20187 | 0.40195 | 0.41284 |  | 0.2402 | 0.3266 | 0.38997 | 0.19408 | 0.52026 | 0.35392 | 0.19792 | 0.31969 | 0.27869 | 0.8165 | 0.25276
| Rel.Std.Dev. |  | 0.8021% | 0.675% | 0.3002% | 2.345% | 2.195% | 3.203% | 3.628% | 2.113% | 1.182% | 2.5% | 2.691% |  | 1.655% | 1.881% | 2.32% | 1.163% | 3.019% | 2.053% | 1.173% | 1.883% | 1.421% | 4.711% | 1.528%
| PDM3 |  | 0.9348% | 6.223% | -1.692% | -11.51% | -0.9417% | 0.1447% | -5.83% | -1.09% | 1.182% | -4.714% | -9.098% |  | -14.01% | 2.91% | -0.3887% | -1.139% | 2.12% | 2.15% | -0.05958% | 0.6187% | 16.24% | 2.712% | -1.983%
| Z-Score (Absolute) |  | 0.38 | 2.12 | 0.48 | 3.71 | 0.24 | 0.12 | 1.84 | 0.28 | 0.46 | 1.48 | 2.92 | NA | 4.53 | 1.03 | 0.05 | 0.3 | 0.77 | 0.78 | 0.05 | 0.28 | 5.41 | 0.97 | 0.58
| Analytical results for Al in OREAS 931b (Certified Value 5.54 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 5.44 | 5.5546 | 5.979 | 5.48 | 5.63 | 5.58 | 5.82 | 5.72 | 5.54 | 3.62 | 5.5777 | 6.06 | 5.38 | 5.4 | 5.2887 | 5.47 | 5.48 | 5.86 | 5.5247 | 5.5 | 5.61 | 5.52 | 5.2864 | 1
| 1 | 2 | 5.39 | 5.4743 | 5.8225 | 5.49 | 5.75 | 5.44 | 5.76 | 5.59 | 5.57 | 3.76 | 5.4637 | 6.07 | 5.55 | 5.7 | 5.3959 | 5.33 | 5.55 | 5.77 | 5.6222 | 5.44 | 5.61 | 5.57 | 5.2967 | 0
| 1 | 3 | 5.51 | 5.6858 | 5.9246 | 5.42 | 5.68 | 5.46 | 5.69 | 5.46 | 5.62 | 3.56 | 5.2389 | 6.16 | 5.42 | 5.7 | 5.1101 | 5.5 | 5.59 | 5.95 | 5.5173 | 5.41 | 5.61 | 5.47 | 5.424 | 16
| 1 | 4 | 5.5 | 5.5039 | 6.0177 | 5.5 | 5.61 | 5.47 | 5.62 | 5.52 | 5.58 | 3.95 | 5.6623 | 6.18 | 5.43 | 5.7 | 5.5628 | 5.29 | 5.94 | 5.87 | 5.5692 | 5.39 | 5.557 | 5.57 | 5.4393 | 5.5432
| 1 | 5 | 5.46 | 5.5425 | 5.893 | 5.51 | 5.7 | 5.48 | 5.5 | 5.6 | 5.53 | 3.83 | 5.3865 | 6.45 | 5.43 | 5.7 | 5.2788 | 5.36 | 5.76 | 5.9 | 5.4336 | 5.53 | 5.663 | 5.58 | 5.2917 | 8
| 1 | 6 | 5.43 | 5.6434 | 5.8512 | 5.52 | 5.72 | 5.47 | 5.58 | 5.61 | 5.56 | 3.42 | 5.3971 | 6.58 | 5.45 | 5.6 | 5.2508 | 5.39 | 5.88 | 5.93 | 5.5788 | 5.45 | 5.716 | 5.5 | 5.4177
| Mean |  | 5.455 | 5.5674 | 5.9147 | 5.4867 | 5.6817 | 5.4833 | 5.6617 | 5.5833 | 5.5667 | 3.69 | 5.4544 | 6.25 | 5.4433 | 5.6333 | 5.3145 | 5.39 | 5.7 | 5.88 | 5.541 | 5.4533 | 5.6277 | 5.535 | 5.3593
| Median |  | 5.45 | 5.5485 | 5.9088 | 5.495 | 5.69 | 5.47 | 5.655 | 5.595 | 5.565 | 3.69 | 5.4304 | 6.17 | 5.43 | 5.7 | 5.2838 | 5.375 | 5.675 | 5.885 | 5.547 | 5.445 | 5.61 | 5.545 | 5.3572
| Std Dev. |  | 0.045056 | 0.081583 | 0.074626 | 0.03559 | 0.053448 | 0.049261 | 0.11839 | 0.088242 | 0.032042 | 0.19329 | 0.15023 | 0.21466 | 0.057155 | 0.12111 | 0.15235 | 0.08124 | 0.18794 | 0.063875 | 0.065107 | 0.053166 | 0.054738 | 0.045056 | 0.074565
| Rel.Std.Dev. |  | 0.8259% | 1.465% | 1.262% | 0.6487% | 0.9407% | 0.8984% | 2.091% | 1.58% | 0.5756% | 5.238% | 2.754% | 3.435% | 1.05% | 2.15% | 2.867% | 1.507% | 3.297% | 1.086% | 1.175% | 0.9749% | 0.9727% | 0.814% | 1.391%
| PDM3 |  | -1.591% | 0.4365% | 6.701% | -1.02% | 2.498% | -1.08% | 2.137% | 0.7239% | 0.4232% | -33.43% | -1.603% | 12.75% | -1.802% | 1.626% | -4.126% | -2.764% | 2.829% | 6.076% | -0.0403% | -1.621% | 1.524% | -0.1481% | -3.318%
| Z-Score (Absolute) |  | 0.63 | 0.19 | 2.73 | 0.4 | 1.03 | 0.42 | 0.88 | 0.31 | 0.19 | 13.51 | 0.63 | 5.18 | 0.71 | 0.67 | 1.65 | 1.1 | 1.16 | 2.48 | 0 | 0.64 | 0.63 | 0.04 | 1.33
| Analytical results for As in OREAS 931b (Certified Value 99 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 0
1
| 1 | 1 | 121 | 101.7845 | 99.8 | 97 | 102.5 | 93.2 | 108 | 105.5 | 101.5 | 93 | 97.4177 | 74 | 95.4 | 97.4 | 97 | 89 | 99.9 | 102.2 | 96.0028 | 90 | 97 | 88.8058 |  | 1
| 1 | 2 | 120 | 99.8781 | 97.8 | 96 | 105.5 | 95.4 | 107.5 | 102 | 103.5 | 100 | 95.3997 | 73 | 95.5 | 100 | 100 | 86 | 101 | 102 | 100.5552 | 92 | 100 | 85.2063 |  | 0
| 1 | 3 | 120 | 102.4454 | 99.3 | 96 | 105.5 | 97.6 | 100 | 101 | 104.5 | 94 | 91.6745 | 75 | 94 | 94.8 | 101 | 89 | 105 | 101.9 | 101.2168 | 93 | 99 | 83.8002 |  | 16
| 1 | 4 | 119 | 101.9555 | 100.3 | 98 | 97.6 | 97.9 | 101.5 | 104 | 104 | 97 | 98.07 | 77 | 95.2 | 102 | 96 | 89 | 105 | 102.2 | 101.0009 | 93 | 100 | 86.5027 |  | 98.5527
| 1 | 5 | 119 | 103.1714 | 98.3 | 97 | 99.5 | 99 | 97.1 | 102 | 105 | 97 | 91.8216 | 77 | 95.1 | 99.3 | 100 | 87 | 107 | 100.7 | 95.0525 | 94 | 106 | 86.6076 |  | 9
| 1 | 6 | 119 | 99.6799 | 97.7 | 97 | 100.5 | 97.3 | 100.5 | 102.5 | 102 | 94 | 94.3053 | 77 | 93.7 | 98.1 | 94 | 90 | 108 | 101.6 | 103.3775 | 94 | 101 | 84.8035
| Mean |  | 119.6667 | 101.4858 | 98.8667 | 96.8333 | 101.85 | 96.7333 | 102.4333 | 102.8333 | 103.4167 | 95.8333 | 94.7815 | 75.5 | 94.8167 | 98.6 | 98 | 88.3333 | 104.3167 | 101.7667 | 99.5343 | 92.6667 | 100.5 | 85.9544
| Median |  | 119.5 | 101.87 | 98.8 | 97 | 101.5 | 97.45 | 101 | 102.25 | 103.75 | 95.5 | 94.8525 | 76 | 95.15 | 98.7 | 98.5 | 89 | 105 | 101.95 | 100.7781 | 93 | 100 | 85.8545
| Std Dev. |  | 0.8165 | 1.4084 | 1.0893 | 0.75277 | 3.2396 | 2.0897 | 4.3734 | 1.633 | 1.3934 | 2.6394 | 2.7125 | 1.7607 | 0.7679 | 2.4552 | 2.7568 | 1.5055 | 3.2314 | 0.56804 | 3.2666 | 1.5055 | 3.0166 | 1.754
| Rel.Std.Dev. |  | 0.6823% | 1.388% | 1.102% | 0.7774% | 3.181% | 2.16% | 4.27% | 1.588% | 1.347% | 2.754% | 2.862% | 2.332% | 0.8099% | 2.49% | 2.813% | 1.704% | 3.098% | 0.5582% | 3.282% | 1.625% | 3.002% | 2.041%
| PDM3 |  | 21.42% | 2.976% | 0.3186% | -1.745% | 3.346% | -1.846% | 3.938% | 4.343% | 4.935% | -2.759% | -3.827% | -23.39% | -3.791% | 0.04798% | -0.5608% | -10.37% | 5.849% | 3.261% | 0.996% | -5.972% | 1.976% | -12.78%
| Z-Score (Absolute) |  | 4.14 | 0.54 | 0.03 | 0.38 | 0.62 | 0.4 | 0.73 | 0.81 | 0.93 | 0.57 | 0.78 | 4.6 | 0.77 | 0.03 | 0.15 | 2.06 | 1.1 | 0.6 | 0.16 | 1.2 | 0.35 | 2.53
| Analytical results for B in OREAS 931b (Indicative Value 51 ppm) | [24 empty] | Indicative
B | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 01 |  | ppm
|  |  | 01: 4A*OES |  | 0
0
| 1 | 1 | 01: 52 |  | 1
| 1 | 2 | 01: 51 |  | 1
| 1 | 3 | 01: 51 |  | 16
| 1 | 4 | 01: 52 |  | 51.3333
| 1 | 5 | 01: 51 |  | 7
| 1 | 6 | 01: 51
| Mean |  | 01: 51.3333
| Median |  | 01: 51
| Std Dev. |  | 01: 0.5164
| Rel.Std.Dev. |  | 01: 1.006%
| PDM3 |  | 01: 6.661e-14%
Z-Score (Absolute) | bar: 01
| Analytical results for Ba in OREAS 931b (Indicative Value 581 ppm) | [24 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 575 | 1649.8196 | 793.2 | 843 | 140 | 140 | 280 | 280 | 370 | 952 | 1921.6542 | 120 | 14 | 79 | 1466 | 152 | 130 | 616 | 666.1 | 205 | 1230 | 119 | 1625.0646 | 1
| 1 | 2 | 591 | 1663.3582 | 941.8 | 825 | 130 | 200 | 290 | 320 | 330 | 746 | 1896.5307 | 94 | 14 | 82 | 1834 | 135 | 22 | 764 | 525 | 201 | 1100 | 79 | 1624.4644 | 3
| 1 | 3 | 570 | 1599.0941 | 769.7 | 813 | 140 | 700 | 270 | 350 | 410 | 587 | 1822.0521 | 102 | 14 | 75 | 1809 | 152 | 25 | 402 | 538 | 178 | 1010 | 78 | 1666.9633 | 16
| 1 | 4 | 563 | 1579.4444 | 838.8 | 808 | 140 | 930 | 340 | 300 | 460 | 964 | 1970.592 | 112 | 13 | 72 | 1729 | 203 | 43 | 455 | 368.8 | 229 | 1080 | 68 | 1658.1638 | 581.1503
| 1 | 5 | 577 | 1601.5234 | 1025.7 | 800 | 140 | 780 | 350 | 350 | 490 | 961 | 1900.898 | 145 | 15 | 80 | 1764 | 414 | 48 | 417 | 424.9 | 217 | 1160 | 62 | 1618.1644 | 7
| 1 | 6 | 576 | 1669.4204 | 839 | 820 | 140 | 900 | 350 | 300 | 490 | 1199 | 1907.4551 | 109 | 15 | 88 | 1726 | 243 | 87 | 374 | 455.9 | 234 | 1040 | 83 | 1650.2636
| Mean |  | 575.3333 | 1627.11 | 868.0333 | 818.1667 | 138.3333 | 608.3333 | 313.3333 | 316.6667 | 425 | 901.5 | 1903.197 | 113.6667 | 14.1667 | 79.3333 | 1721.3333 | 216.5 | 59.1667 | 504.6667 | 496.45 | 210.6667 | 1103.3333 | 81.5 | 1640.514
| Median |  | 575.5 | 1625.6715 | 838.9 | 816.5 | 140 | 740 | 315 | 310 | 435 | 956.5 | 1904.1765 | 110.5 | 14 | 79.5 | 1746.5 | 177.5 | 45.5 | 436 | 490.45 | 211 | 1090 | 78.5 | 1637.6641
| Std Dev. |  | 9.2664 | 38.2921 | 97.2289 | 15.0122 | 4.0825 | 349.9952 | 37.238 | 28.7518 | 66.2571 | 210.5201 | 48.0315 | 17.7163 | 0.75277 | 5.5737 | 132.2689 | 104.7411 | 41.7776 | 153.2419 | 104.2889 | 20.5589 | 80.6639 | 19.9474 | 20.5032
| Rel.Std.Dev. |  | 1.611% | 2.353% | 11.2% | 1.835% | 2.951% | 57.53% | 11.88% | 9.08% | 15.59% | 23.35% | 2.524% | 15.59% | 5.314% | 7.026% | 7.684% | 48.38% | 70.61% | 30.36% | 21.01% | 9.759% | 7.311% | 24.48% | 1.25%
| PDM3 |  | -1.001% | 180% | 49.36% | 40.78% | -76.2% | 4.677% | -46.08% | -45.51% | -26.87% | 55.12% | 227.5% | -80.44% | -97.56% | -86.35% | 196.2% | -62.75% | -89.82% | -13.16% | -14.57% | -63.75% | 89.85% | -85.98% | 182.3%
| Z-Score (Absolute) |  | 0.14 | 2.05 | 0.67 | 0.58 | 0.65 | 0.2 | 0.33 | 0.33 | 0.13 | 0.74 | 2.55 | 0.69 | 0.88 | 0.76 | 2.22 | 0.51 | 0.79 | 0.01 | 0 | 0.52 | 1.1 | 0.75 | 2.08
| Analytical results for Be in OREAS 931b (Certified Value 2.08 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | < 5 | 2.1724 | 2.31 | 1.6 | 1.96 | 2.08 | 2.07 | 2.06 | 1.98 | 2.5 | 2.1648 | 2.75 | 2 | 1.94 | 1.84 | 2 | 1.9 | 2.16 | 2.2853 | 1.84 | 2.2 | 1.9 | 1.7875 | 1
| 1 | 2 | < 5 | 2.1765 | 2.16 | 1.7 | 1.99 | 2.09 | 2.05 | 2 | 2 | 2.7 | 2.1556 | 2.82 | 3 | 1.99 | 1.85 | 2 | 1.9 | 2.15 | 2.3626 | 2.18 | 2.1 | 2.5 | 1.7935 | 0
| 1 | 3 | < 5 | 2.1746 | 2.25 | 1.7 | 1.96 | 2 | 2.04 | 1.99 | 2.02 | 2.4 | 2.0409 | 3.16 | 3 | 2.06 | 1.77 | 3 | 2 | 2.2 | 2.3681 | 2.05 | 2.1 | 1.9 | 1.8474 | 16
| 1 | 4 | < 5 | 2.1839 | 2.17 | 1.7 | 1.96 | 1.99 | 2.01 | 1.99 | 1.99 | 2.5 | 2.2633 | 3.22 | 3 | 1.95 | 1.89 | 3 | 2.1 | 2.17 | 2.4477 | 2.04 | 2.2 | 1.9 | 1.8328 | 2.0818
| 1 | 5 | < 5 | 2.1787 | 2.28 | 1.8 | 2 | 2 | 1.95 | 2.02 | 2 | 2.5 | 2.1681 | 3.04 | 3 | 1.93 | 1.75 | 2 | 2.1 | 2.19 | 2.4387 | 2.04 | 2.2 | 2.5 | 1.7927 | 10
| 1 | 6 | < 5 | 2.1808 | 2.16 | 1.8 | 1.98 | 2.01 | 1.98 | 2 | 1.99 | 2.3 | 2.2139 | 2.99 | 3 | 1.91 | 1.79 | 2 | 2 | 2.21 | 2.5269 | 2.12 | 2.2 | 1.8 | 1.8275
| Mean |  |  | 2.1778 | 2.2217 | 1.7167 | 1.975 | 2.0283 | 2.0167 | 2.01 | 1.9967 | 2.4833 | 2.1678 | 2.9967 | 2.8333 | 1.9633 | 1.815 | 2.3333 | 2 | 2.18 | 2.4049 | 2.045 | 2.1667 | 2.0833 | 1.8136
| Median |  |  | 2.1776 | 2.21 | 1.7 | 1.97 | 2.005 | 2.025 | 2 | 1.995 | 2.5 | 2.1665 | 3.015 | 3 | 1.945 | 1.815 | 2 | 2 | 2.18 | 2.4034 | 2.045 | 2.2 | 1.9 | 1.8105
| Std Dev. |  |  | 0.0042199 | 0.066758 | 0.075277 | 0.017607 | 0.04446 | 0.045461 | 0.026833 | 0.013663 | 0.13292 | 0.074171 | 0.18468 | 0.40825 | 0.054283 | 0.053572 | 0.5164 | 0.089443 | 0.023664 | 0.084026 | 0.11485 | 0.05164 | 0.32506 | 0.025405
| Rel.Std.Dev. |  |  | 0.1938% | 3.005% | 4.385% | 0.8915% | 2.192% | 2.254% | 1.335% | 0.6843% | 5.352% | 3.422% | 6.163% | 14.41% | 2.765% | 2.952% | 22.13% | 4.472% | 1.086% | 3.494% | 5.616% | 2.383% | 15.6% | 1.401%
| PDM3 |  |  | 4.615% | 6.721% | -17.54% | -5.128% | -2.566% | -3.127% | -3.447% | -4.088% | 19.29% | 4.131% | 43.95% | 36.1% | -5.689% | -12.81% | 12.08% | -3.928% | 4.719% | 15.52% | -1.766% | 4.079% | 0.07551% | -12.88%
| Z-Score (Absolute) |  | 2.31 | 0.67 | 0.9 | NA | 0.36 | 0.08 | 0.14 | 0.18 | 0.25 | 2.23 | 0.62 | 4.83 | NA | 0.41 | 1.17 | NA | NA | 0.69 | 1.83 | 0 | NA | NA | 1.17
Results from laboratories 4, 17, 22 and 23 were removed due to their 0.1 ppm reading resolution. Results from laboratories 13 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 931b (Certified Value 211 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 197 | 231.2148 | 211.6 | 199 | 222 | 234 | 221 | 225 | 201 | 190.62 | 202.5351 | 198 | 221.18 | 221 | 210.85 | 225.3 | 205 | 193.45 | 221.1076 | 201 | 240 | 177.9417 |  | 1
| 1 | 2 | 196 | 230.3466 | 218.87 | 194 | 217 | 229 | 225 | 220 | 205 | 199.09 | 204.3493 | 198 | 222.34 | 224 | 215.85 | 229.5 | 207 | 183.89 | 222.0419 | 215 | 254 | 181.9422 |  | 8
| 1 | 3 | 206 | 233.9726 | 217.19 | 192 | 223 | 210 | 218 | 214 | 203 | 183.62 | 198.2151 | 209 | 218.99 | 216 | 207.09 | 231.6 | 212 | 180.87 | 221.7943 | 211 | 242 | 187.9393 |  | 16
| 1 | 4 | 198 | 233.1431 | 221.55 | 196 | 221 | 212 | 220 | 225 | 203 | 196.08 | 215.6203 | 212 | 217.74 | 225 | 219.53 | 223.1 | 192 | 185.98 | 219 | 206 | 226 | 188.9407 |  | 211.3386
| 1 | 5 | 207 | 232.3502 | 219.24 | 193 | 224 | 212 | 207 | 218 | 198 | 198.42 | 208.8859 | 211 | 216.69 | 220 | 214.68 | 226 | 193 | 192.64 | 218.8323 | 212 | 237 | 186.9427 |  | 11
| 1 | 6 | 201 | 231.3448 | 211.93 | 188 | 222 | 215 | 214 | 217 | 202 | 190.93 | 207.7103 | 214 | 216.04 | 218 | 208.01 | 232.1 | 194 | 181.19 | 218.833 | 207 | 239 | 184.9405
| Mean |  | 200.8333 | 232.062 | 216.73 | 193.6667 | 221.5 | 218.6667 | 217.5 | 219.8333 | 202 | 193.1267 | 206.2193 | 207 | 218.83 | 220.6667 | 212.6683 | 227.9333 | 200.5 | 186.3367 | 220.2682 | 208.6667 | 239.6667 | 184.7745
| Median |  | 199.5 | 231.8475 | 218.03 | 193.5 | 222 | 213.5 | 219 | 219 | 202.5 | 193.505 | 206.0298 | 210 | 218.365 | 220.5 | 212.765 | 227.75 | 199.5 | 184.935 | 220.0538 | 209 | 239.5 | 185.9416
| Std Dev. |  | 4.7081 | 1.3474 | 4.091 | 3.7238 | 2.429 | 10.1915 | 6.2849 | 4.446 | 2.3664 | 5.8956 | 5.9808 | 7.1554 | 2.5068 | 3.4448 | 4.8451 | 3.6686 | 8.5499 | 5.5284 | 1.5433 | 5.0067 | 9.0037 | 4.1668
| Rel.Std.Dev. |  | 2.344% | 0.5806% | 1.888% | 1.923% | 1.097% | 4.661% | 2.89% | 2.022% | 1.172% | 3.053% | 2.9% | 3.457% | 1.146% | 1.561% | 2.278% | 1.61% | 4.264% | 2.967% | 0.7007% | 2.399% | 3.757% | 2.255%
| PDM3 |  | -4.971% | 9.806% | 2.551% | -8.362% | 4.808% | 3.467% | 2.915% | 4.019% | -4.419% | -8.617% | -2.422% | -2.053% | 3.545% | 4.414% | 0.6292% | 7.852% | -5.129% | -11.83% | 4.225% | -1.264% | 13.4% | -12.57%
| Z-Score (Absolute) |  | 1.16 | 1.45 | 0.17 | 1.75 | 0.57 | 0.33 | 0.23 | 0.43 | 1.06 | 1.8 | 0.71 | 0.64 | 0.34 | 0.5 | 0.17 | 1.1 | 1.18 | 2.36 | 0.46 | 0.5 | 2.08 | 2.49
| Analytical results for Ca in OREAS 931b (Certified Value 0.083 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.08 | 0.085644 | 0.0872 | 0.081 | 0.08 | 0.08 | 0.09 | 0.09 | 0.08 | 0.07 | 0.107 | 0.09 | 0.08 | 0.074 | 0.08 | 0.08 | 0.09 | 0.08838 | 0.08 | 0.086 | 0.09 | 0.10034 |  | 1
| 1 | 2 | 0.08 | 0.083905 | 0.0858 | 0.081 | 0.09 | 0.08 | 0.09 | 0.09 | 0.08 | 0.07 | 0.114 | 0.09 | 0.08 | 0.0792 | 0.08 | 0.08 | 0.08 | 0.09182 | 0.079 | 0.086 | 0.08 | 0.10113 |  | 0
| 1 | 3 | 0.08 | 0.0848 | 0.0859 | 0.082 | 0.08 | 0.08 | 0.09 | 0.08 | 0.08 | 0.07 | 0.12 | 0.09 | 0.08 | 0.0745 | 0.08 | 0.08 | 0.09 | 0.09062 | 0.079 | 0.086 | 0.08 | 0.10363 |  | 16
| 1 | 4 | 0.08 | 0.086313 | 0.0867 | 0.081 | 0.08 | 0.08 | 0.09 | 0.08 | 0.08 | 0.07 | 0.126 | 0.09 | 0.08 | 0.0778 | 0.08 | 0.09 | 0.09 | 0.0888 | 0.079 | 0.086 | 0.08 | 0.10353 |  | 0.083318
| 1 | 5 | 0.09 | 0.08377 | 0.0854 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.06 | 0.144 | 0.09 | 0.08 | 0.0775 | 0.08 | 0.08 | 0.09 | 0.08827 | 0.08 | 0.086 | 0.09 | 0.10064 |  | 12
| 1 | 6 | 0.08 | 0.086354 | 0.0864 | 0.081 | 0.08 | 0.08 | 0.08 | 0.09 | 0.08 | 0.06 | 0.133 | 0.09 | 0.08 | 0.0764 | 0.08 | 0.08 | 0.09 | 0.08814 | 0.08 | 0.086 | 0.08 | 0.10243
| Mean |  | 0.081667 | 0.085131 | 0.086233 | 0.081 | 0.081667 | 0.08 | 0.086667 | 0.086667 | 0.08 | 0.066667 | 0.124 | 0.09 | 0.08 | 0.076567 | 0.08 | 0.081667 | 0.088333 | 0.089338 | 0.0795 | 0.086 | 0.083333 | 0.10195
| Median |  | 0.08 | 0.085222 | 0.08615 | 0.081 | 0.08 | 0.08 | 0.09 | 0.09 | 0.08 | 0.07 | 0.123 | 0.09 | 0.08 | 0.07695 | 0.08 | 0.08 | 0.09 | 0.08859 | 0.0795 | 0.086 | 0.08 | 0.10178
| Std Dev. |  | 0.0040825 | 0.0011508 | 0.00065929 | 0.00063246 | 0.0040825 | 0 | 0.005164 | 0.005164 | 0 | 0.005164 | 0.013342 | 1.5202e-17 | 0 | 0.0020106 | 0 | 0.0040825 | 0.0040825 | 0.0015223 | 0.00054772 | 1.5202e-17 | 0.005164 | 0.0014527
| Rel.Std.Dev. |  | 4.999% | 1.352% | 0.7645% | 0.7808% | 4.999% | 0% | 5.958% | 5.958% | 0% | 7.746% | 10.76% | 1.689e-14% | 0% | 2.626% | 0% | 4.999% | 4.622% | 1.704% | 0.689% | 1.768e-14% | 6.197% | 1.425%
| PDM3 |  | -1.982% | 2.177% | 3.499% | -2.782% | -1.982% | -3.982% | 4.019% | 4.019% | -3.982% | -19.99% | 48.83% | 8.02% | -3.982% | -8.103% | -3.982% | -1.982% | 6.02% | 7.226% | -4.582% | 3.219% | 0.01875% | 22.36%
| Z-Score (Absolute) |  | 0.18 | 0.57 | 0.8 | 0.32 | 0.18 | 0.54 | 0.9 | 0.9 | 0.54 | 3.41 | 8.94 | 1.61 | 0.54 | 1.28 | 0.54 | 0.18 | 1.26 | 1.47 | 0.65 | 0.7 | 0.18 | 4.19
| Analytical results for Cd in OREAS 931b (Certified Value 0.64 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 2
2
| 1 | 1 | 2 | 0.65814 | 0.71 | 0.61 | 0.42 | 0.53 | 0.52 | 0.57 | 0.72 | 0.62 | < 5 | 0.53 | 0.7 | 0.66 | 0.64 | 0.7 | 0.6 | 0.66 | 0.66609 | 0.65 | < 1 | < 1 | 0.6121 | 1
| 1 | 2 | 2 | 0.66511 | 0.66 | 0.63 | 0.44 | 0.57 | 0.46 | 0.58 | 0.78 | 0.63 | < 5 | 0.48 | 0.74 | 0.69 | 0.63 | 0.8 | 0.6 | 0.66 | 0.68137 | 0.64 | < 1 | 1 | 0.6798 | 20
| 1 | 3 | 2 | 0.66154 | 0.58 | 0.62 | 0.51 | 0.62 | 0.46 | 0.54 | 0.76 | 0.64 | < 5 | 0.49 | 0.75 | 0.64 | 0.57 | 0.5 | 0.6 | 0.68 | 0.64793 | 0.75 | < 1 | < 1 | 0.5679 | 16
| 1 | 4 | 1.9 | 0.65804 | 0.65 | 0.62 | 0.4 | 0.64 | 0.44 | 0.56 | 0.79 | 0.7 | < 5 | 0.59 | 0.73 | 0.68 | 0.57 | 0.6 | 0.6 | 0.67 | 0.68296 | 0.62 | < 1 | 1 | 0.6827 | 0.64385
| 1 | 5 | 2 | 0.64522 | 0.69 | 0.66 | 0.44 | 0.66 | 0.42 | 0.56 | 0.77 | 0.61 | < 5 | 0.57 | 0.73 | 0.67 | 0.6 | 0.8 | 0.8 | 0.63 | 0.66238 | 0.67 | < 1 | < 1 | 0.6328 | 13
| 1 | 6 | 2 | 0.64593 | 0.69 | 0.62 | 0.45 | 0.53 | 0.41 | 0.61 | 0.79 | 0.64 | < 5 | 0.54 | 0.75 | 0.64 | 0.61 | 0.7 | 0.5 | 0.65 | 0.68454 | 0.67 | < 1 | < 1 | 0.614
| Mean |  | 1.9833 | 0.65566 | 0.66333 | 0.62667 | 0.44333 | 0.59167 | 0.45167 | 0.57 | 0.76833 | 0.64 |  | 0.53333 | 0.73333 | 0.66333 | 0.60333 | 0.68333 | 0.61667 | 0.65833 | 0.67088 | 0.66667 |  | 1 | 0.63155
| Median |  | 2 | 0.65809 | 0.675 | 0.62 | 0.44 | 0.595 | 0.45 | 0.565 | 0.775 | 0.635 |  | 0.535 | 0.735 | 0.665 | 0.605 | 0.7 | 0.6 | 0.66 | 0.67373 | 0.66 |  | 1 | 0.6234
| Std Dev. |  | 0.040825 | 0.0082379 | 0.046332 | 0.017512 | 0.037238 | 0.056362 | 0.0392 | 0.023664 | 0.026394 | 0.031623 |  | 0.043205 | 0.018619 | 0.020656 | 0.029439 | 0.1169 | 0.098319 | 0.017224 | 0.014589 | 0.045019 |  | 0 | 0.044008
| Rel.Std.Dev. |  | 2.058% | 1.256% | 6.985% | 2.794% | 8.4% | 9.526% | 8.679% | 4.152% | 3.435% | 4.941% |  | 8.101% | 2.539% | 3.114% | 4.879% | 17.11% | 15.94% | 2.616% | 2.175% | 6.753% |  | 0% | 6.968%
| PDM3 |  | 208% | 1.835% | 3.026% | -2.669% | -31.14% | -8.105% | -29.85% | -11.47% | 19.33% | -0.5981% |  | -17.17% | 13.9% | 3.026% | -6.293% | 6.132% | -4.222% | 2.249% | 4.198% | 3.544% |  | 55.32% | -1.91%
| Z-Score (Absolute) |  | NA | 0.16 | 0.31 | 0.42 | 4.08 | 1.12 | 3.92 | 1.55 | 2.41 | 0.16 | 36.98 | 2.29 | 1.71 | 0.31 | 0.89 | NA | NA | 0.21 | 0.46 | 0.38 | 2.95 | 0.38 | 0.33
| Analytical results for Ce in OREAS 931b (Certified Value 71 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 24 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 76.06 | 70.9 | 69.8 | 74.1 | 68.6 | 69.6 | 62 | 77.9453 | 82.38 | 62.79 | 72.8 | 75.2 | 64 | 62 | 70.36 | 70.63 | 78 | 56.8604 | [5 empty] | 1
| 1 | 2 | 77.5 | 71.2 | 72.3 | 74.8 | 68.9 | 70.9 | 66 | 75.7593 | 85.7 | 63.03 | 73.3 | 77.68 | 67 | 65.4 | 67.13 | 76.53 | 76 | 64.5203 | [5 empty] | 9
| 1 | 3 | 76.85 | 75 | 72.4 | 69.1 | 65.1 | 72.7 | 61.05 | 74.0161 | 94.31 | 63.79 | 71.8 | 75.35 | 66 | 68.4 | 72.16 | 73.59 | 80 | 58.8627 | [5 empty] | 16
| 1 | 4 | 77.71 | 70.5 | 70.4 | 70.2 | 69.5 | 72.8 | 67.63 | 75.9501 | 101 | 66.42 | 75.8 | 80.37 | 65 | 68.4 | 70.41 | 74.23 | 79 | 66.6884 | [5 empty] | 70.6165
| 1 | 5 | 77.73 | 75.2 | 70.1 | 69.3 | 68.4 | 69.7 | 66.16 | 74.3712 | 107 | 63.41 | 77.2 | 78.53 | 66 | 70.4 | 69.61 | 74.65 | 80 | 52.4017 | [5 empty] | 14
| 1 | 6 | 76.4 | 75.4 | 71 | 68.1 | 69.5 | 73.1 | 61.68 | 73.0322 | 99.61 | 64.27 | 75.2 | 76.91 | 65 | 71.4 | 68.39 | 74.02 | 82 | 58.5373
| Mean |  | 77.0417 | 73.0333 | 71 | 70.9333 | 68.3333 | 71.4667 | 64.0867 | 75.179 | 95 | 63.9517 | 74.35 | 77.34 | 65.5 | 67.6667 | 69.6767 | 73.9417 | 79.1667 | 59.6451
| Median |  | 77.175 | 73.1 | 70.7 | 69.75 | 68.75 | 71.8 | 64 | 75.0653 | 96.96 | 63.6 | 74.25 | 77.295 | 65.5 | 68.4 | 69.985 | 74.125 | 79.5 | 58.7
| Std Dev. |  | 0.71312 | 2.3872 | 1.1189 | 2.8133 | 1.6476 | 1.6058 | 2.8243 | 1.7434 | 9.4603 | 1.32 | 2.0472 | 1.9717 | 1.0488 | 3.4564 | 1.7503 | 1.9174 | 2.0412 | 5.2045
| Rel.Std.Dev. |  | 0.9256% | 3.269% | 1.576% | 3.966% | 2.411% | 2.247% | 4.407% | 2.319% | 9.958% | 2.064% | 2.753% | 2.549% | 1.601% | 5.108% | 2.512% | 2.593% | 2.578% | 8.726%
| PDM3 |  | 9.099% | 3.422% | 0.5431% | 0.4487% | -3.233% | 1.204% | -9.247% | 6.461% | 34.53% | -9.438% | 5.287% | 9.521% | -7.245% | -4.177% | -1.331% | 4.709% | 12.11% | -15.54%
| Z-Score (Absolute) |  | 1.17 | 0.36 | 0.05 | 0.06 | 0.59 | 0.05 | 1.44 | 0.8 | 4.8 | 1.47 | 0.63 | 1.23 | NA | 0.72 | 0.31 | 0.55 | NA | 2.34
Results from laboratories 16 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 931b (Certified Value 29.6 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 30 | 30.7756 | 30.2 | 30 | 29.7 | 26.6 | 30.5 | 29.1 | 27.8 | 29.9 | 28.5531 | 28.9 | 29.3 | 30.4 | 29 | 29.4 | 30 | 30.6 | 29.7967 | 24.3 | 16 | 5 | 25.3489 | 1
| 1 | 2 | 32 | 30.5849 | 29.9 | 31 | 29.6 | 26.1 | 30.3 | 28.6 | 28.3 | 31.6 | 28.1176 | 28.8 | 29.4 | 30.6 | 29.8 | 29.9 | 30 | 31.6 | 30.1203 | 25.6 | 19 | 6 | 25.1492 | 11
| 1 | 3 | 32 | 30.6994 | 30.1 | 31 | 30.1 | 27.5 | 29.2 | 28.1 | 28.8 | 29.3 | 26.9184 | 30.2 | 28.8 | 29.2 | 28.4 | 30.6 | 30 | 31.7 | 30.3518 | 26.4 | 17 | 5 | 25.8107 | 16
| 1 | 4 | 32 | 30.6289 | 30.5 | 31 | 28.3 | 27.2 | 28.2 | 29.4 | 28.3 | 30.6 | 29.2063 | 31.4 | 29 | 31.1 | 30.1 | 29.5 | 31 | 30.1 | 30.4531 | 26.5 | 16 | 6 | 25.9241 | 29.5646
| 1 | 5 | 32 | 30.7235 | 29.8 | 31 | 29.3 | 28.3 | 28.3 | 28.1 | 28.9 | 30.7 | 27.9212 | 30.2 | 29.1 | 30.9 | 29.2 | 30.1 | 31 | 31.3 | 30.1489 | 26.4 | 18 | 6 | 25.6437 | 15
| 1 | 6 | 31 | 30.6997 | 29.7 | 31 | 29.6 | 27.6 | 27.9 | 28.6 | 28.3 | 29.4 | 27.7353 | 30.7 | 29 | 30.1 | 29.2 | 31.1 | 31 | 31.1 | 30.7815 | 26 | 18 | 6 | 25.7144
| Mean |  | 31.5 | 30.6853 | 30.0333 | 30.8333 | 29.4333 | 27.2167 | 29.0667 | 28.65 | 28.4 | 30.25 | 28.0753 | 30.0333 | 29.1 | 30.3833 | 29.2833 | 30.1 | 30.5 | 31.0667 | 30.2754 | 25.8667 | 17.3333 | 5.6667 | 25.5985
| Median |  | 32 | 30.6995 | 30 | 31 | 29.6 | 27.35 | 28.75 | 28.6 | 28.3 | 30.25 | 28.0194 | 30.2 | 29.05 | 30.5 | 29.2 | 30 | 30.5 | 31.2 | 30.2504 | 26.2 | 17.5 | 6 | 25.679
| Std Dev. |  | 0.83666 | 0.068226 | 0.29439 | 0.40825 | 0.6121 | 0.77825 | 1.1219 | 0.5244 | 0.4 | 0.88261 | 0.77292 | 1.0172 | 0.21909 | 0.67946 | 0.60139 | 0.65422 | 0.54772 | 0.61536 | 0.33548 | 0.83825 | 1.2111 | 0.5164 | 0.29347
| Rel.Std.Dev. |  | 2.656% | 0.2223% | 0.9802% | 1.324% | 2.08% | 2.859% | 3.86% | 1.83% | 1.408% | 2.918% | 2.753% | 3.387% | 0.7529% | 2.236% | 2.054% | 2.173% | 1.796% | 1.981% | 1.108% | 3.241% | 6.987% | 9.113% | 1.146%
| PDM3 |  | 6.546% | 3.791% | 1.585% | 4.291% | -0.4441% | -7.942% | -1.684% | -3.094% | -3.939% | 2.318% | -5.037% | 1.585% | -1.572% | 2.769% | -0.9515% | 1.811% | 3.164% | 5.081% | 2.404% | -12.51% | -41.37% | -80.83% | -13.42%
| Z-Score (Absolute) |  | 1.35 | 0.82 | 0.39 | 0.91 | 0 | 1.45 | 0.24 | 0.51 | 0.67 | 0.53 | 0.89 | 0.39 | 0.22 | 0.62 | 0.1 | 0.44 | 0.7 | 1.07 | 0.55 | 2.33 | 7.9 | 15.51 | 2.5
| Analytical results for Cr in OREAS 931b (Certified Value 72 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS |  | 0
1
| 1 | 1 | 79 | 72.0373 | 87 | 81 | 72 | 72 | 66 | 73 | 70 | 53 | 88 | 76 | 74.2 | 74 | 59 | 111 | 67 | 66.682 | 90 | 91 | 69 | 71.6268 |  | 1
| 1 | 2 | 76 | 72.9949 | 85 | 80 | 73 | 71 | 68 | 71 | 71 | 56 | 90 | 66 | 74.7 | 74 | 61 | 70 | 64 | 65.5769 | 88 | 90 | 77 | 71.4255 |  | 0
| 1 | 3 | 77 | 72.6475 | 87 | 79 | 73 | 66 | 73 | 73 | 71 | 49 | 91 | 64 | 72.6 | 75 | 61 | 63 | 67 | 75.3028 | 87 | 91 | 76 | 73.9245 |  | 16
| 1 | 4 | 76 | 72.5219 | 89 | 79 | 73 | 65 | 71 | 68 | 72 | 53 | 91 | 63 | 72.1 | 72 | 63 | 61 | 72 | 83.1074 | 84 | 92 | 77 | 72.6237 |  | 72.1313
| 1 | 5 | 75 | 73.6912 | 86 | 77 | 72 | 66 | 73 | 69 | 72 | 54 | 96 | 68 | 69.8 | 74 | 66 | 86 | 68 | 73.4276 | 89 | 89 | 84 | 71.3266 |  | 16
| 1 | 6 | 77 | 73.4759 | 84 | 78 | 72 | 65 | 71 | 72 | 71 | 56 | 97 | 65 | 70.2 | 70 | 64 | 73 | 58 | 70.8601 | 88 | 91 | 73 | 72.7246
| Mean |  | 76.6667 | 72.8948 | 86.3333 | 79 | 72.5 | 67.5 | 70.3333 | 71 | 71.1667 | 53.5 | 92.1667 | 67 | 72.2667 | 73.1667 | 62.3333 | 77.3333 | 66 | 72.4928 | 87.6667 | 90.6667 | 76 | 72.2753
| Median |  | 76.5 | 72.8212 | 86.5 | 79 | 72.5 | 66 | 71 | 71.5 | 71 | 53.5 | 91 | 65.5 | 72.35 | 74 | 62 | 71.5 | 67 | 72.1439 | 88 | 91 | 76.5 | 72.1252
| Std Dev. |  | 1.3663 | 0.61925 | 1.7512 | 1.4142 | 0.54772 | 3.1464 | 2.8048 | 2.0976 | 0.75277 | 2.5884 | 3.5449 | 4.7329 | 2.0077 | 1.8348 | 2.5033 | 18.7261 | 4.6904 | 6.414 | 2.0656 | 1.0328 | 4.98 | 1.0086
| Rel.Std.Dev. |  | 1.782% | 0.8495% | 2.028% | 1.79% | 0.7555% | 4.661% | 3.988% | 2.954% | 1.058% | 4.838% | 3.846% | 7.064% | 2.778% | 2.508% | 4.016% | 24.21% | 7.107% | 8.848% | 2.356% | 1.139% | 6.553% | 1.395%
| PDM3 |  | 6.288% | 1.059% | 19.69% | 9.523% | 0.5112% | -6.421% | -2.493% | -1.568% | -1.337% | -25.83% | 27.78% | -7.114% | 0.1877% | 1.435% | -13.58% | 7.212% | -8.5% | 0.5012% | 21.54% | 25.7% | 5.363% | 0.1997%
| Z-Score (Absolute) |  | 0.62 | 0.06 | 2.07 | 0.97 | 0 | 0.75 | 0.32 | 0.22 | 0.2 | 2.84 | 2.95 | 0.82 | 0.03 | 0.1 | 1.52 | 0.72 | 0.97 | 0 | 2.27 | 2.72 | 0.52 | 0.03
| Analytical results for Cs in OREAS 931b (Certified Value 4.75 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 5.0522 | 4.87 | 4.4 | 4.3 | 4.68 | 4.91 | 4.76 | 4.37 | < 5 | 4.7556 | 5.14 | 4.3 | 4.88 | 4.63 | 4.74 | 4.86 | 5.2288 | 4.33 | 5.0748 |  |  |  |  | 1
| 1 | 2 | 5.0222 | 4.72 | 4.5 | 4.35 | 4.82 | 4.87 | 4.76 | 4.42 | < 5 | 4.6731 | 5.18 | 4.4 | 4.94 | 4.75 | 4.8 | 4.92 | 5.2597 | 4.64 | 5.0811 |  |  |  |  | 0
| 1 | 3 | 5.0523 | 4.8 | 4.5 | 4.51 | 4.73 | 4.59 | 4.47 | 4.5 | < 5 | 4.4307 | 5.54 | 4.4 | 4.78 | 4.5 | 4.88 | 4.97 | 5.2719 | 4.62 | 5.1432 |  |  |  |  | 16
| 1 | 4 | 5.0404 | 4.88 | 4.4 | 4.37 | 4.57 | 4.64 | 4.76 | 4.41 | < 5 | 4.8644 | 5.66 | 4.3 | 4.98 | 4.8 | 4.89 | 4.85 | 5.2707 | 4.54 | 5.173 |  |  |  |  | 4.7453
| 1 | 5 | 4.9807 | 4.8 | 4.7 | 4.49 | 4.71 | 4.55 | 4.59 | 4.53 | < 5 | 4.7542 | 5.76 | 4.3 | 4.9 | 4.7 | 5.12 | 4.98 | 5.2348 | 4.54 | 4.8707 |  |  |  |  | 17
| 1 | 6 | 4.9851 | 4.81 | 4.5 | 4.44 | 4.66 | 4.53 | 4.79 | 4.57 | < 5 | 4.7035 | 5.82 | 4.4 | 4.8 | 4.62 | 5.17 | 5 | 5.3017 | 4.51 | 4.9426
| Mean |  | 5.0222 | 4.8133 | 4.5 | 4.41 | 4.695 | 4.6817 | 4.6883 | 4.4667 |  | 4.6969 | 5.5167 | 4.35 | 4.88 | 4.6667 | 4.9333 | 4.93 | 5.2613 | 4.53 | 5.0476
| Median |  | 5.0313 | 4.805 | 4.5 | 4.405 | 4.695 | 4.615 | 4.76 | 4.46 |  | 4.7289 | 5.6 | 4.35 | 4.89 | 4.665 | 4.885 | 4.945 | 5.2652 | 4.54 | 5.078
| Std Dev. |  | 0.032362 | 0.05785 | 0.10954 | 0.083187 | 0.082644 | 0.16618 | 0.12891 | 0.078145 |  | 0.14578 | 0.29242 | 0.054772 | 0.077974 | 0.1069 | 0.17363 | 0.063875 | 0.026802 | 0.11027 | 0.11751
| Rel.Std.Dev. |  | 0.6444% | 1.202% | 2.434% | 1.886% | 1.76% | 3.55% | 2.749% | 1.75% |  | 3.104% | 5.301% | 1.259% | 1.598% | 2.291% | 3.519% | 1.296% | 0.5094% | 2.434% | 2.328%
| PDM3 |  | 5.835% | 1.434% | -5.169% | -7.066% | -1.06% | -1.341% | -1.2% | -5.871% |  | -1.019% | 16.26% | -8.33% | 2.839% | -1.657% | 3.963% | 3.893% | 10.87% | -4.537% | 6.37%
| Z-Score (Absolute) |  | 0.97 | 0.35 | 0.58 | 0.84 | 0 | 0.04 | 0.02 | 0.67 | 6.48 | 0.01 | 2.43 | 1.02 | 0.55 | 0.08 | 0.7 | 0.69 | 1.67 | 0.49 | 1.04
| Analytical results for Cu in OREAS 931b (Certified Value 3.95 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 4.13 | 4.1354 | 3.952 | 3.96 | 3.84 | 3.94 | 3.9 | 3.84 | 3.88 | 4.17 | 3.9453 | 4.1457 | 3.869 | 3.83 | 3.8954 | 4.109 | 4.05 | 4.111 | 3.825 | 3.88 | 4.11 | 3.615 | 3.6883 | 1
| 1 | 2 | 4.15 | 4.0965 | 4.0818 | 3.97 | 3.87 | 3.87 | 3.92 | 3.88 | 3.87 | 4.14 | 3.878 | 4.22 | 3.923 | 3.81 | 3.8618 | 4.028 | 4.01 | 3.997 | 3.899 | 4.03 | 4.12 | 3.863 | 3.727 | 0
| 1 | 3 | 4.16 | 4.0031 | 4.0622 | 3.93 | 3.87 | 3.84 | 3.99 | 3.91 | 3.88 | 4.03 | 3.7212 | 4.2858 | 3.973 | 3.81 | 3.8846 | 4.02 | 4.07 | 3.875 | 3.82 | 4.05 | 4.09 | 3.895 | 3.8028 | 16
| 1 | 4 | 4.18 | 3.9806 | 4.0799 | 3.98 | 3.87 | 3.89 | 3.79 | 3.95 | 3.82 | 4.16 | 4.0291 | 4.3163 | 3.954 | 3.84 | 3.8443 | 4.111 | 4.07 | 3.99 | 3.864 | 3.95 | 4.08 | 3.785 | 3.7897 | 3.9471
| 1 | 5 | 4.14 | 3.9932 | 4.1568 | 3.91 | 3.87 | 3.87 | 3.89 | 3.81 | 3.84 | 4.09 | 3.8616 | 4.5461 | 3.967 | 3.83 | 3.8554 | 3.954 | 4.11 | 4.039 | 3.756 | 4.01 | 4.06 | 3.533 | 3.7081 | 18
| 1 | 6 | 4.17 | 4.1288 | 4.0678 | 3.88 | 3.87 | 3.81 | 3.93 | 3.89 | 3.86 | 4.16 | 3.8342 | 4.6422 | 3.94 | 3.82 | 3.9017 | 4.078 | 4.13 | 4.043 | 3.856 | 4.02 | 4.07 | 3.45 | 3.766
| Mean |  | 4.155 | 4.0563 | 4.0667 | 3.9383 | 3.865 | 3.87 | 3.9033 | 3.88 | 3.8583 | 4.125 | 3.8782 | 4.3594 | 3.9377 | 3.8233 | 3.8739 | 4.05 | 4.0733 | 4.0092 | 3.8367 | 3.99 | 4.0883 | 3.6902 | 3.747
| Median |  | 4.155 | 4.0498 | 4.0739 | 3.945 | 3.87 | 3.87 | 3.91 | 3.885 | 3.865 | 4.15 | 3.8698 | 4.3011 | 3.947 | 3.825 | 3.8732 | 4.053 | 4.07 | 4.018 | 3.8405 | 4.015 | 4.085 | 3.7 | 3.7465
| Std Dev. |  | 0.018708 | 0.071653 | 0.065875 | 0.038687 | 0.012247 | 0.044272 | 0.065625 | 0.0498 | 0.024014 | 0.054681 | 0.10408 | 0.19348 | 0.03824 | 0.012111 | 0.0233 | 0.060969 | 0.04274 | 0.078639 | 0.048841 | 0.063561 | 0.023166 | 0.18377 | 0.046166
| Rel.Std.Dev. |  | 0.4503% | 1.766% | 1.62% | 0.9823% | 0.3169% | 1.144% | 1.681% | 1.283% | 0.6224% | 1.326% | 2.684% | 4.438% | 0.9711% | 0.3168% | 0.6015% | 1.505% | 1.049% | 1.961% | 1.273% | 1.593% | 0.5666% | 4.98% | 1.232%
| PDM3 |  | 5.266% | 2.764% | 3.03% | -0.2232% | -2.081% | -1.954% | -1.11% | -1.701% | -2.25% | 4.506% | -1.745% | 10.44% | -0.24% | -3.137% | -1.856% | 2.606% | 3.197% | 1.571% | -2.799% | 1.086% | 3.577% | -6.51% | -5.071%
| Z-Score (Absolute) |  | 1.45 | 0.79 | 0.86 | 0 | 0.49 | 0.45 | 0.23 | 0.39 | 0.53 | 1.25 | 0.4 | 2.82 | 0 | 0.76 | 0.43 | 0.75 | 0.91 | 0.48 | 0.67 | 0.35 | 1.01 | 1.65 | 1.27
| Analytical results for Dy in OREAS 931b (Certified Value 3.06 ppm) | [24 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 03; 02: 11; 03: 12; 04: 13; 05: 15; 06: 17; 07: 22; 08: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES; 08: 4A*OES/MS | 2
2
| 1 | 1 | 01: 2.9; 02: 2.7135; 03: 2.91; 04: 3; 05: 2.76; 06: 3.5; 07: 3.3; 08: 1.8289 | 1
| 1 | 2 | 01: 2.86; 02: 2.6783; 03: 3.04; 04: 3; 05: 2.94; 06: 3.5; 07: 3.2; 08: 1.8743 | 4
| 1 | 3 | 01: 2.86; 02: 2.6233; 03: 3.11; 04: 3; 05: 2.8; 06: 3.4; 07: 3.2; 08: 1.8212 | 16
| 1 | 4 | 01: 2.84; 02: 2.804; 03: 3.38; 04: 2.8; 05: 2.89; 06: 3.7; 07: 3.3; 08: 1.8451 | 3.0608
| 1 | 5 | 01: 2.88; 02: 2.7549; 03: 3.02; 04: 3; 05: 2.83; 06: 3.9; 07: 3.3; 08: 1.9164 | 19
| 1 | 6 | 01: 2.85; 02: 2.7102; 03: 3.44; 04: 2.9; 05: 2.86; 06: 3.9; 07: 3.2; 08: 1.9531
| Mean |  | 01: 2.865; 02: 2.714; 03: 3.15; 04: 2.95; 05: 2.8467; 06: 3.65; 07: 3.25; 08: 1.8732
| Median |  | 01: 2.86; 02: 2.7119; 03: 3.075; 04: 3; 05: 2.845; 06: 3.6; 07: 3.25; 08: 1.8597
| Std Dev. |  | 01: 0.021679; 02: 0.06211; 03: 0.21223; 04: 0.083666; 05: 0.064395; 06: 0.21679; 07: 0.054772; 08: 0.052359
| Rel.Std.Dev. |  | 01: 0.7567%; 02: 2.288%; 03: 6.737%; 04: 2.836%; 05: 2.262%; 06: 5.94%; 07: 1.685%; 08: 2.795%
| PDM3 |  | 01: -6.397%; 02: -11.33%; 03: 2.914%; 04: -3.62%; 05: -6.996%; 06: 19.25%; 07: 6.181%; 08: -38.8%
| Z-Score (Absolute) |  | 01: 0.13; 02: 0.6; 03: 0.75; 04: 0.13; 05: 0.19; 06: 2.3; 07: 1.06; 08: 3.2
| Analytical results for Er in OREAS 931b (Certified Value 1.37 ppm) | [24 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 03; 02: 11; 03: 12; 04: 13; 05: 15; 06: 17; 07: 22; 08: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES; 08: 4A*OES/MS | 2
3
| 1 | 1 | 01: 1.41; 02: 1.3371; 03: 1.31; 04: 1.5; 05: 1.3; 06: 2.1; 07: < 0.2; 08: 1.2498 | 1
| 1 | 2 | 01: 1.39; 02: 1.3302; 03: 1.26; 04: 1.5; 05: 1.35; 06: 2; 07: < 0.2; 08: 1.2486 | 5
| 1 | 3 | 01: 1.34; 02: 1.3063; 03: 1.42; 04: 1.6; 05: 1.3; 06: 2.1; 07: < 0.2; 08: 1.2842 | 16
| 1 | 4 | 01: 1.4; 02: 1.3683; 03: 1.49; 04: 1.4; 05: 1.35; 06: 2.1; 07: < 0.2; 08: 1.2996 | 1.3743
| 1 | 5 | 01: 1.37; 02: 1.353; 03: 1.49; 04: 1.5; 05: 1.39; 06: 2.2; 07: < 0.2; 08: 1.2512 | 20
| 1 | 6 | 01: 1.38; 02: 1.3264; 03: 1.56; 04: 1.5; 05: 1.32; 06: 2.2; 07: < 0.2; 08: 1.2886
| Mean |  | 01: 1.3817; 02: 1.3369; 03: 1.4217; 04: 1.5; 05: 1.335; 06: 2.1167; 08: 1.2703
| Median |  | 01: 1.385; 02: 1.3337; 03: 1.455; 04: 1.5; 05: 1.335; 06: 2.1; 08: 1.2677
| Std Dev. |  | 01: 0.024833; 02: 0.021618; 03: 0.11583; 04: 0.063246; 05: 0.035071; 06: 0.075277; 08: 0.022996
| Rel.Std.Dev. |  | 01: 1.797%; 02: 1.617%; 03: 8.148%; 04: 4.216%; 05: 2.627%; 06: 3.556%; 08: 1.81%
| PDM3 |  | 01: 0.5391%; 02: -2.72%; 03: 3.45%; 04: 9.15%; 05: -2.857%; 06: 54.02%; 08: -7.562%
| Z-Score (Absolute) |  | 01: 0.2; 02: 0.2; 03: 0.56; 04: 1.25; 05: 0.22; 06: 6.75; 07: 11.22; 08: 0.79
| Analytical results for Eu in OREAS 931b (Certified Value 1.09 ppm) | [24 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 03; 02: 12; 03: 13; 04: 15; 05: 17; 06: 22; 07: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES; 07: 4A*OES/MS | 2
3
| 1 | 1 | 01: 1.08; 02: 1.07; 03: 1.2; 04: 1.07; 05: 1.08; 06: 1.7; 07: 0.9054 | 1
| 1 | 2 | 01: 1.07; 02: 1.1; 03: 1.2; 04: 1.11; 05: 1.07; 06: 1.7; 07: 0.91069 | 6
| 1 | 3 | 01: 1.1; 02: 1.27; 03: 1.2; 04: 1.05; 05: 1.08; 06: 1.7; 07: 0.93276 | 16
| 1 | 4 | 01: 1.08; 02: 1.33; 03: 1.1; 04: 1.11; 05: 1.05; 06: 1.7; 07: 0.92157 | 1.0906
| 1 | 5 | 01: 1.09; 02: 1.27; 03: 1.1; 04: 1.08; 05: 1.16; 06: 1.8; 07: 0.90603 | 21
| 1 | 6 | 01: 1.07; 02: 1.29; 03: 1.1; 04: 1.06; 05: 1.12; 06: 1.7; 07: 0.92645
| Mean |  | 01: 1.0817; 02: 1.2217; 03: 1.15; 04: 1.08; 05: 1.0933; 06: 1.7167; 07: 0.91715
| Median |  | 01: 1.08; 02: 1.27; 03: 1.15; 04: 1.075; 05: 1.08; 06: 1.7; 07: 0.91613
| Std Dev. |  | 01: 0.01169; 02: 0.10852; 03: 0.054772; 04: 0.025298; 05: 0.039833; 06: 0.040825; 07: 0.01143
| Rel.Std.Dev. |  | 01: 1.081%; 02: 8.883%; 03: 4.763%; 04: 2.342%; 05: 3.643%; 06: 2.378%; 07: 1.246%
| PDM3 |  | 01: -0.8224%; 02: 12.01%; 03: 5.443%; 04: -0.9752%; 05: 0.2473%; 06: 57.4%; 07: -15.91%
| Z-Score (Absolute) |  | 01: 0.14; 02: 1.53; 03: 0.67; 04: 0.16; 05: 0; 06: 7.42; 07: 2.1
| Analytical results for Fe in OREAS 931b (Certified Value 8.55 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 8.6 | 8.7246 | 9.06 | 8.2 | 8.57 | 8.58 | 8.78 | 8.64 | 8.35 | 8 | 8.8178 | 9.26 | 8.25 | 8.71 | 7.9679 | 8.35 | 8.63 | 8.98 | 9.144 | 8.24 | 8.953 | 8.2 | 8.5521 | 1
| 1 | 2 | 8.59 | 8.6111 | 8.94 | 8.17 | 8.78 | 8.36 | 8.75 | 8.47 | 8.42 | 8.4 | 8.6275 | 9.42 | 8.51 | 9.11 | 8.3748 | 8.34 | 8.65 | 8.7 | 9.5798 | 8.15 | 8.953 | 8.27 | 8.6175 | 0
| 1 | 3 | 8.63 | 8.9218 | 9.07 | 8.08 | 8.65 | 8.17 | 8.63 | 8.23 | 8.54 | 7.85 | 8.18 | 9.52 | 8.26 | 8.64 | 7.9258 | 8.48 | 8.99 | 9.13 | 9.2294 | 8.05 | 8.953 | 8.26 | 8.8793 | 16
| 1 | 4 | 8.62 | 8.6558 | 9.15 | 8.31 | 8.56 | 8.15 | 8.55 | 8.33 | 8.44 | 8.18 | 8.816 | 9.58 | 8.32 | 8.88 | 8.3561 | 8.17 | 8.53 | 9 | 9.216 | 8.07 | 8.883 | 8.27 | 8.831 | 8.552
| 1 | 5 | 8.59 | 8.7199 | 8.83 | 8.04 | 8.71 | 8.21 | 8.3 | 8.46 | 8.38 | 8.23 | 8.4166 | 10.05 | 8.38 | 8.84 | 8.1761 | 8.34 | 8.68 | 9 | 9.2506 | 8.2 | 9.023 | 8.7 | 8.5473 | 22
| 1 | 6 | 8.65 | 8.8859 | 8.92 | 8.02 | 8.72 | 8.17 | 8.46 | 8.46 | 8.4 | 7.73 | 8.4007 | 10.29 | 8.34 | 8.62 | 8.24 | 8.37 | 8.98 | 9.14 | 9.2592 | 8.17 | 9.023 | 8.24 | 8.7574
| Mean |  | 8.6133 | 8.7532 | 8.995 | 8.1367 | 8.665 | 8.2733 | 8.5783 | 8.4317 | 8.4217 | 8.065 | 8.5431 | 9.6867 | 8.3433 | 8.8 | 8.1735 | 8.3417 | 8.7433 | 8.9917 | 9.2798 | 8.1467 | 8.9647 | 8.3233 | 8.6974
| Median |  | 8.61 | 8.7222 | 9 | 8.125 | 8.68 | 8.19 | 8.59 | 8.46 | 8.41 | 8.09 | 8.5221 | 9.55 | 8.33 | 8.775 | 8.2081 | 8.345 | 8.665 | 9 | 9.24 | 8.16 | 8.953 | 8.265 | 8.6875
| Std Dev. |  | 0.024221 | 0.12462 | 0.11811 | 0.11075 | 0.087807 | 0.1686 | 0.1817 | 0.13963 | 0.065853 | 0.25114 | 0.255 | 0.39707 | 0.095219 | 0.1845 | 0.19075 | 0.099482 | 0.19387 | 0.15905 | 0.15255 | 0.073937 | 0.052694 | 0.1864 | 0.1446
| Rel.Std.Dev. |  | 0.2812% | 1.424% | 1.313% | 1.361% | 1.013% | 2.038% | 2.118% | 1.656% | 0.782% | 3.114% | 2.985% | 4.099% | 1.141% | 2.097% | 2.334% | 1.193% | 2.217% | 1.769% | 1.644% | 0.9076% | 0.5878% | 2.24% | 1.663%
| PDM3 |  | 0.7176% | 2.353% | 5.18% | -4.856% | 1.322% | -3.258% | 0.3083% | -1.407% | -1.524% | -5.694% | -0.1036% | 13.27% | -2.44% | 2.9% | -4.426% | -2.459% | 2.238% | 5.142% | 8.511% | -4.739% | 4.826% | -2.673% | 1.701%
| Z-Score (Absolute) |  | 0.1 | 0.5 | 1.19 | 1.26 | 0.25 | 0.87 | 0 | 0.42 | 0.45 | 1.46 | 0.1 | 3.16 | 0.67 | 0.63 | 1.15 | 0.67 | 0.47 | 1.18 | 2 | 1.23 | 1.1 | 0.73 | 0.34
| Analytical results for Ga in OREAS 931b (Certified Value 17.7 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 24 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 22 | 18.3424 | 18.44 | 20.2 | 16.75 | 16.35 | 16.8 | 18.25 | 17.9 | 16.9 | 18.5579 | 16.38 | 18.36 | 18.9 | 16.82 | 17.3 | 18.58 | 18.9749 | 15.18 | 26 | 16.409 |  |  | 1
| 1 | 2 | 22 | 18.4248 | 17.56 | 20.8 | 16.8 | 16.65 | 16.65 | 17.85 | 18 | 17.8 | 18.1771 | 15.86 | 18.34 | 18.9 | 17.79 | 17.8 | 17.95 | 17.9514 | 16.13 | 25 | 16.8403 |  |  | 14
| 1 | 3 | 22 | 18.9232 | 17.92 | 20.8 | 17.35 | 17.6 | 15.6 | 17.7 | 18 | 16.6 | 17.2529 | 17.37 | 18.12 | 17.9 | 16.53 | 18.5 | 18.71 | 18.0612 | 16.19 | 25 | 16.0421 |  |  | 16
| 1 | 4 | 22 | 19.0591 | 18.46 | 21.1 | 16.3 | 17.15 | 15.75 | 18.7 | 18.6 | 17.4 | 18.5734 | 17.27 | 18.07 | 19.3 | 17.55 | 18.9 | 18.24 | 17.9867 | 15.95 | 25 | 16.1889 |  |  | 17.7095
| 1 | 5 | 22 | 19.2069 | 18.13 | 20.5 | 16.6 | 17.3 | 15.85 | 17.65 | 17.75 | 17.5 | 17.7743 | 16.95 | 18.33 | 19 | 17.04 | 19.7 | 18.72 | 17.9726 | 16.62 | 26 | 14.9024 |  |  | 23
| 1 | 6 | 22 | 18.0127 | 17.81 | 20.1 | 16.7 | 17.4 | 15.3 | 18.15 | 18.6 | 17 | 17.8428 | 17.52 | 17.42 | 18.5 | 16.91 | 19 | 17.41 | 18.2689 | 15.61 | 25 | 15.5338
| Mean |  | 22 | 18.6615 | 18.0533 | 20.5833 | 16.75 | 17.075 | 15.9917 | 18.05 | 18.1417 | 17.2 | 18.0297 | 16.8917 | 18.1067 | 18.75 | 17.1067 | 18.5333 | 18.2683 | 18.2026 | 15.9467 | 25.3333 | 15.9861
| Median |  | 22 | 18.674 | 18.025 | 20.65 | 16.725 | 17.225 | 15.8 | 18 | 18 | 17.2 | 18.0099 | 17.11 | 18.225 | 18.9 | 16.975 | 18.7 | 18.41 | 18.0239 | 16.04 | 25 | 16.1155
| Std Dev. |  | 0 | 0.46966 | 0.35809 | 0.38687 | 0.34351 | 0.47828 | 0.59951 | 0.39875 | 0.36663 | 0.44272 | 0.50988 | 0.6475 | 0.3582 | 0.48888 | 0.47357 | 0.86872 | 0.51674 | 0.39583 | 0.49947 | 0.5164 | 0.6829
| Rel.Std.Dev. |  | 0% | 2.517% | 1.983% | 1.88% | 2.051% | 2.801% | 3.749% | 2.209% | 2.021% | 2.574% | 2.828% | 3.833% | 1.978% | 2.607% | 2.768% | 4.687% | 2.829% | 2.175% | 3.132% | 2.038% | 4.272%
| PDM3 |  | 24.23% | 5.376% | 1.941% | 16.23% | -5.418% | -3.583% | -9.7% | 1.922% | 2.44% | -2.877% | 1.808% | -4.618% | 2.242% | 5.875% | -3.404% | 4.652% | 3.155% | 2.784% | -9.954% | 43.05% | -9.732%
| Z-Score (Absolute) |  | 3.44 | 0.53 | 0 | 2.2 | 1.13 | 0.85 | 1.79 | 0 | 0.08 | 0.74 | 0.02 | 1.01 | 0.05 | 0.61 | 0.82 | 0.42 | 0.19 | 0.13 | 1.83 | NA | 1.8
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 931b (Certified Value 4.43 ppm) | [24 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  |  | 1
| No | No | 01: 03; 02: 12; 03: 13; 04: 15; 05: 17; 06: 22; 07: 24 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*OES; 07: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 01: 4.39; 02: 4.8; 03: 4.2; 04: 4.18; 05: 4.9; 06: < 1; 07: 3.9101 |  |  | 1
| 1 | 2 | 01: 4.33; 02: 4.69; 03: 4.2; 04: 4.35; 05: 5.2; 06: < 1; 07: 3.9921 |  |  | 15
| 1 | 3 | 01: 4.35; 02: 5.09; 03: 4.4; 04: 4.17; 05: 5.2; 06: < 1; 07: 3.9188 |  |  | 16
| 1 | 4 | 01: 4.32; 02: 5.13; 03: 4; 04: 4.42; 05: 4.4; 06: < 1; 07: 4.1285 |  |  | 4.4278
| 1 | 5 | 01: 4.33; 02: 4.89; 03: 4.1; 04: 4.28; 05: 4.8; 06: < 1; 07: 3.851 |  |  | 24
| 1 | 6 | 01: 4.29; 02: 5.34; 03: 4.1; 04: 4.2; 05: 4.6; 06: < 1; 07: 3.9497
| Mean |  | 01: 4.335; 02: 4.99; 03: 4.1667; 04: 4.2667; 05: 4.85; 07: 3.9584
| Median |  | 01: 4.33; 02: 4.99; 03: 4.15; 04: 4.24; 05: 4.85; 07: 3.9343
| Std Dev. |  | 01: 0.033317; 02: 0.24008; 03: 0.13663; 04: 0.10191; 05: 0.32094; 07: 0.095479
| Rel.Std.Dev. |  | 01: 0.7686%; 02: 4.811%; 03: 3.279%; 04: 2.389%; 05: 6.617%; 07: 2.412%
| PDM3 |  | 01: -2.095%; 02: 12.7%; 03: -5.897%; 04: -3.639%; 05: 9.536%; 07: -10.6%
| Z-Score (Absolute) |  | 01: 0.15; 02: 1.58; 03: 0.22; 04: 0; 05: 1.28; 06: 8.24; 07: 0.67
| Analytical results for Ge in OREAS 931b (Certified Value 0.21 ppm) | [24 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 15 | 17 | 18 | 19 | [8 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 2
2
| 1 | 1 | 1.9621 | 2.1 | 1.8 | 0.24 | 0.22 | 0.21 | 0.16 | 0.22 | 0.8 | 0.55 | 0.2 | 1.6 | 0.3 | 0.17 | 0.60507 | [8 empty] | 1
| 1 | 2 | 1.9649 | 2.1 | 1.8 | 0.24 | 0.23 | 0.19 | 0.21 | 0.23 | 0.8 | 0.44 | 0.2 | 1.6 | 0.3 | 0.17 | 0.64484 | [8 empty] | 0
| 1 | 3 | 2.0174 | 2 | 1.8 | 0.24 | 0.28 | 0.17 | 0.16 | 0.26 | 0.8 | 0.46 | 0.16 | 1.5 | 0.3 | 0.19 | 0.70658 | [8 empty] | 16
| 1 | 4 | 1.9941 | 2.2 | 1.8 | 0.22 | 0.29 | 0.17 | 0.17 | 0.22 | 0.8 | 0.55 | 0.2 | 1.6 | 0.3 | 0.18 | 0.82271 | [8 empty] | 0.20952
| 1 | 5 | 2.0344 | 2.1 | 1.9 | 0.23 | 0.29 | 0.17 | 0.16 | 0.24 | 0.8 | 0.44 | 0.2 | 1.7 | 0.3 | 0.18 | 0.82485 | [8 empty] | 25
| 1 | 6 | 1.9245 | 2.3 | 1.7 | 0.22 | 0.29 | 0.18 | 0.2 | 0.23 | 0.8 | 0.56 | 0.2 | 1.7 | 0.3 | 0.17 | 0.97904
| Mean |  | 1.9829 | 2.1333 | 1.8 | 0.23167 | 0.26667 | 0.18167 | 0.17667 | 0.23333 | 0.8 | 0.5 | 0.19333 | 1.6167 | 0.3 | 0.17667 | 0.76385
| Median |  | 1.9795 | 2.1 | 1.8 | 0.235 | 0.285 | 0.175 | 0.165 | 0.23 | 0.8 | 0.505 | 0.2 | 1.6 | 0.3 | 0.175 | 0.76464
| Std Dev. |  | 0.040336 | 0.10328 | 0.063246 | 0.0098319 | 0.03266 | 0.016021 | 0.022509 | 0.015055 | 1.2162e-16 | 0.058992 | 0.01633 | 0.075277 | 0 | 0.008165 | 0.13867
| Rel.Std.Dev. |  | 2.034% | 4.841% | 3.514% | 4.244% | 12.25% | 8.819% | 12.74% | 6.452% | 1.52e-14% | 11.8% | 8.447% | 4.656% | 0% | 4.622% | 18.15%
| PDM3 |  | 846.4% | 918.2% | 759.1% | 10.57% | 27.27% | -13.3% | -15.68% | 11.36% | 281.8% | 138.6% | -7.727% | 671.6% | 43.18% | -15.68% | 264.6%
| Z-Score (Absolute) |  | 22.13 | NA | NA | 0.01 | 0.43 | 0.64 | 0.71 | 0.01 | NA | 3.38 | 0.5 | NA | NA | 0.71 | 6.72
Results from laboratories 3, 4, 10, 15 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 931b (Certified Value 1.82 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 1.8352 | 1.79 | 1.8 | 1.8 | 1.8 | 1.9 | 2 | 1.8 | 1.81 | 1.94 | 1.79 | 1.92 | 1.71 | 1.8 | 2 | 1.9 | 1.4415 | 1.83 | 17 | 1.6259 |  |  |  | 1
| 1 | 2 | 1.8251 | 1.8 | 1.8 | 1.8 | 1.9 | 1.9 | 1.9 | 1.7 | 1.86 | 1.94 | 1.79 | 1.93 | 1.82 | 1.8 | 2 | 1.8 | 1.3479 | 1.91 | 16 | 1.6506 |  |  |  | 16
| 1 | 3 | 1.8204 | 1.84 | 1.8 | 1.7 | 1.9 | 1.8 | 1.9 | 1.8 | 1.76 | 2.1 | 1.85 | 1.87 | 1.69 | 1.8 | 2 | 2 | 1.3892 | 1.89 | 15 | 1.7296 |  |  |  | 16
| 1 | 4 | 1.8224 | 1.85 | 1.8 | 1.8 | 1.8 | 1.8 | 1.9 | 1.7 | 1.81 | 2.22 | 1.74 | 1.94 | 1.79 | 1.7 | 2 | 1.9 | 1.3101 | 1.86 | 16 | 1.7152 |  |  |  | 1.8155
| 1 | 5 | 1.8239 | 1.84 | 1.9 | 1.8 | 1.8 | 1.8 | 1.8 | 1.8 | 1.86 | 2.22 | 1.72 | 1.95 | 1.75 | 1.6 | 2 | 2 | 1.3621 | 1.87 | 15 | 1.6333 |  |  |  | 26
| 1 | 6 | 1.8376 | 1.82 | 1.8 | 1.8 | 1.9 | 1.8 | 1.9 | 1.9 | 1.82 | 2.26 | 1.73 | 1.89 | 1.73 | 1.7 | 2 | 1.8 | 1.365 | 1.87 | 16 | 1.6716
| Mean |  | 1.8274 | 1.8233 | 1.8167 | 1.7833 | 1.85 | 1.8333 | 1.9 | 1.7833 | 1.82 | 2.1133 | 1.77 | 1.9167 | 1.7483 | 1.7333 | 2 | 1.9 | 1.3693 | 1.8717 | 15.8333 | 1.671
| Median |  | 1.8245 | 1.83 | 1.8 | 1.8 | 1.85 | 1.8 | 1.9 | 1.8 | 1.815 | 2.16 | 1.765 | 1.925 | 1.74 | 1.75 | 2 | 1.9 | 1.3635 | 1.87 | 16 | 1.6611
| Std Dev. |  | 0.0071694 | 0.024221 | 0.040825 | 0.040825 | 0.054772 | 0.05164 | 0.063246 | 0.075277 | 0.037417 | 0.14459 | 0.049396 | 0.030768 | 0.04916 | 0.08165 | 0 | 0.089443 | 0.043917 | 0.027142 | 0.75277 | 0.043034
| Rel.Std.Dev. |  | 0.3923% | 1.328% | 2.247% | 2.289% | 2.961% | 2.817% | 3.329% | 4.221% | 2.056% | 6.842% | 2.791% | 1.605% | 2.812% | 4.711% | 0% | 4.708% | 3.207% | 1.45% | 4.754% | 2.575%
| PDM3 |  | 0.6555% | 0.4299% | 0.06266% | -1.773% | 1.899% | 0.9807% | 4.653% | -1.773% | 0.2463% | 16.4% | -2.508% | 5.571% | -3.701% | -4.527% | 10.16% | 4.653% | -24.58% | 3.092% | 772.1% | -7.959%
| Z-Score (Absolute) |  | 0.08 | 0.02 | 0.07 | 0.51 | 0.38 | 0.15 | 1.04 | 0.51 | 0.02 | 3.87 | 0.69 | 1.26 | 0.97 | 1.17 | NA | 1.04 | 6 | 0.66 | NA | 2
Results from laboratory 17 were removed due to their 0.1 ppm reading resolution. Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Hg in OREAS 931b (Indicative Value 0.16 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | 02: Lab |  |  |  | 1
| No | No | 02: 24 |  |  |  | ppm
|  |  | 02: 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 02: 0.1567 |  |  |  | 1
| 1 | 2 | 02: 0.1573 |  |  |  | 1
| 1 | 3 | 02: 0.1562 |  |  |  | 16
| 1 | 4 | 02: 0.1573 |  |  |  | 0.15515
| 1 | 5 | 02: 0.1505 |  |  |  | 7
| 1 | 6 | 02: 0.1529
| Mean |  | 02: 0.15515
| Median |  | 02: 0.15645
| Std Dev. |  | 02: 0.0028084
| Rel.Std.Dev. |  | 02: 1.81%
| PDM3 |  | 02: -2.22e-14%
Z-Score (Absolute) | bar: 02
| Analytical results for Ho in OREAS 931b (Certified Value 0.49 ppm) | [24 empty] | Certified
Ho | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  |  | 1
| No | No | 02: 03; 03: 11; 04: 12; 05: 13; 06: 15; 07: 17; 08: 19; 09: 22; 10: 24 |  |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*OES; 10: 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 02: 0.5; 03: 0.5027; 04: 0.45; 05: 0.6; 06: 0.47; 07: 0.7; 08: 0.62437; 09: < 1; 10: 0.4061 |  |  |  | 1
| 1 | 2 | 02: 0.5; 03: 0.4889; 04: 0.48; 05: 0.5; 06: 0.5; 07: 0.7; 08: 0.58297; 09: < 1; 10: 0.4073 |  |  |  | 4
| 1 | 3 | 02: 0.49; 03: 0.489; 04: 0.54; 05: 0.5; 06: 0.48; 07: 0.7; 08: 0.57873; 09: < 1; 10: 0.419 |  |  |  | 16
| 1 | 4 | 02: 0.51; 03: 0.4985; 04: 0.53; 05: 0.5; 06: 0.5; 07: 0.7; 08: 0.54579; 09: < 1; 10: 0.4257 |  |  |  | 0.48863
| 1 | 5 | 02: 0.5; 03: 0.5024; 04: 0.53; 05: 0.5; 06: 0.49; 07: 0.8; 08: 0.65959; 09: < 1; 10: 0.4126 |  |  |  | 27
| 1 | 6 | 02: 0.49; 03: 0.4964; 04: 0.59; 05: 0.5; 06: 0.48; 07: 0.8; 08: 0.65513; 09: < 1; 10: 0.412
| Mean |  | 02: 0.49833; 03: 0.49632; 04: 0.52; 05: 0.51667; 06: 0.48667; 07: 0.73333; 08: 0.60776; 10: 0.41378
| Median |  | 02: 0.5; 03: 0.49745; 04: 0.53; 05: 0.5; 06: 0.485; 07: 0.7; 08: 0.60367; 10: 0.4123
| Std Dev. |  | 02: 0.0075277; 03: 0.0061817; 04: 0.04899; 05: 0.040825; 06: 0.012111; 07: 0.05164; 08: 0.045835; 10: 0.0074144
| Rel.Std.Dev. |  | 02: 1.511%; 03: 1.246%; 04: 9.421%; 05: 7.902%; 06: 2.488%; 07: 7.042%; 08: 7.542%; 10: 1.792%
| PDM3 |  | 02: 1.986%; 03: 1.574%; 04: 6.42%; 05: 5.738%; 06: -0.4014%; 07: 50.08%; 08: 24.38%; 10: -15.32%
| Z-Score (Absolute) |  | 02: 0.07; 03: 0.15; 04: 0.81; 05: 0.67; 06: 0.54; 07: 9.44; 08: 4.36; 09: 0; 10: 3.49
| Analytical results for In in OREAS 931b (Certified Value 3.32 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | < 2 | 3.4598 | 3.22 | 3.46 | 3.46 | 3.2 | 3.45 | 3.49 | 3.36 | 3.03 | 3.3478 | 2.354 | 3.23 | 3.3 | 3.21 | 3.77 | 3.3 | 3.239 | 3.3697 | 3.12 | 3.4731 |  |  | 1
| 1 | 2 | < 2 | 3.4522 | 3.21 | 3.47 | 3.43 | 3.36 | 3.53 | 3.41 | 3.37 | 3.19 | 3.29 | 2.44 | 3.19 | 3.32 | 3.28 | 3.19 | 3.1 | 3.167 | 3.4276 | 3.233 | 3.5612 |  |  | 0
| 1 | 3 | < 2 | 3.4619 | 3.28 | 3.46 | 3.49 | 3.14 | 3.26 | 3.27 | 3.49 | 2.96 | 3.1119 | 2.556 | 3.26 | 3.21 | 3.08 | 3.59 | 3.1 | 3.212 | 3.3868 | 3.176 | 3.548 |  |  | 16
| 1 | 4 | 2 | 3.3979 | 3.28 | 3.47 | 3.34 | 3.11 | 3.29 | 3.52 | 3.38 | 3.17 | 3.3795 | 2.616 | 3.12 | 3.34 | 3.23 | 3.47 | 3.3 | 3.221 | 3.4204 | 3.133 | 3.57 |  |  | 3.3152
| 1 | 5 | < 2 | 3.433 | 3.2 | 3.51 | 3.51 | 3.11 | 3.23 | 3.43 | 3.41 | 3.23 | 3.4328 | 2.568 | 3.2 | 3.33 | 3.16 | 3.42 | 3.5 | 3.297 | 3.4204 | 3.224 | 3.3789 |  |  | 28
| 1 | 6 | < 2 | 3.3885 | 3.24 | 3.4 | 3.45 | 3.1 | 3.19 | 3.52 | 3.41 | 3.07 | 3.2445 | 2.618 | 3.13 | 3.24 | 3.05 | 3.59 | 3.4 | 3.311 | 3.4286 | 3.155 | 3.4408
| Mean |  | 2 | 3.4322 | 3.2383 | 3.4617 | 3.4467 | 3.17 | 3.325 | 3.44 | 3.4033 | 3.1083 | 3.3011 | 2.5253 | 3.1883 | 3.29 | 3.1683 | 3.505 | 3.2833 | 3.2412 | 3.4089 | 3.1735 | 3.4953
| Median |  | 2 | 3.4426 | 3.23 | 3.465 | 3.455 | 3.125 | 3.275 | 3.46 | 3.395 | 3.12 | 3.3189 | 2.562 | 3.195 | 3.31 | 3.185 | 3.53 | 3.3 | 3.23 | 3.4204 | 3.1655 | 3.5105
| Std Dev. |  |  | 0.032021 | 0.034881 | 0.035449 | 0.059554 | 0.1 | 0.13442 | 0.095079 | 0.047188 | 0.10477 | 0.11384 | 0.10602 | 0.054924 | 0.052915 | 0.089312 | 0.19614 | 0.16021 | 0.054326 | 0.024599 | 0.046779 | 0.077087
| Rel.Std.Dev. |  |  | 0.933% | 1.077% | 1.024% | 1.728% | 3.155% | 4.043% | 2.764% | 1.387% | 3.371% | 3.449% | 4.198% | 1.723% | 1.608% | 2.819% | 5.596% | 4.879% | 1.676% | 0.7216% | 1.474% | 2.205%
| PDM3 |  | -39.67% | 3.528% | -2.32% | 4.417% | 3.964% | -4.381% | 0.2944% | 3.763% | 2.657% | -6.241% | -0.427% | -23.83% | -3.828% | -0.7613% | -4.431% | 5.724% | -0.9624% | -2.234% | 2.826% | -4.275% | 5.432%
| Z-Score (Absolute) |  | NA | 0.74 | 0.31 | 0.9 | 0.82 | 0.68 | 0.16 | 0.79 | 0.59 | 1.02 | 0.03 | 4.2 | 0.58 | 0.03 | 0.69 | 1.14 | 0.07 | 0.3 | 0.62 | 0.66 | 1.09
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for K in OREAS 931b (Certified Value 2.18 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.17 | 2.1337 | 2.3454 | 2.06 | 2.17 | 2.14 | 2.27 | 2.21 | 2.14 | 2.14 | 2.2109 | 2.36 | 2.24 | 2.1 | 2.1584 | 2.12 | 2.17 | 2.26 | 2.294 | 2.07 | 2.183 | 2.32 | 2.1236 | 1
| 1 | 2 | 2.14 | 2.1044 | 2.3071 | 2.07 | 2.21 | 2.06 | 2.24 | 2.16 | 2.15 | 2.23 | 2.173 | 2.4 | 2.25 | 2.12 | 2.1739 | 2.11 | 2.23 | 2.2 | 2.335 | 2.07 | 2.183 | 2.33 | 2.1227 | 0
| 1 | 3 | 2.16 | 2.1885 | 2.3311 | 2.08 | 2.16 | 2.09 | 2.25 | 2.13 | 2.17 | 2.08 | 2.0753 | 2.41 | 2.23 | 2.11 | 2.204 | 2.15 | 2.3 | 2.29 | 2.309 | 2.04 | 2.142 | 2.3 | 2.199 | 16
| 1 | 4 | 2.17 | 2.1154 | 2.3484 | 2.08 | 2.17 | 2.09 | 2.2 | 2.14 | 2.15 | 2.17 | 2.2344 | 2.43 | 2.24 | 2.09 | 2.1264 | 2.06 | 2.01 | 2.28 | 2.333 | 2.02 | 2.167 | 2.34 | 2.1899 | 2.1761
| 1 | 5 | 2.15 | 2.1303 | 2.3117 | 2.06 | 2.18 | 2.11 | 2.13 | 2.17 | 2.14 | 2.18 | 2.1341 | 2.57 | 2.25 | 2.03 | 2.1831 | 2.09 | 2.25 | 2.28 | 2.25 | 2.07 | 2.192 | 2.14 | 2.119 | 29
| 1 | 6 | 2.15 | 2.1747 | 2.2724 | 2.07 | 2.19 | 2.11 | 2.16 | 2.17 | 2.14 | 2.07 | 2.1327 | 2.62 | 2.22 | 2.04 | 2.0559 | 2.1 | 2.19 | 2.31 | 2.323 | 2.05 | 2.192 | 2.32 | 2.1704
| Mean |  | 2.1567 | 2.1412 | 2.3194 | 2.07 | 2.18 | 2.1 | 2.2083 | 2.1633 | 2.1483 | 2.145 | 2.1601 | 2.465 | 2.2383 | 2.0817 | 2.1503 | 2.105 | 2.1917 | 2.27 | 2.3073 | 2.0533 | 2.1765 | 2.2917 | 2.1541
| Median |  | 2.155 | 2.132 | 2.3214 | 2.07 | 2.175 | 2.1 | 2.22 | 2.165 | 2.145 | 2.155 | 2.1536 | 2.42 | 2.24 | 2.095 | 2.1662 | 2.105 | 2.21 | 2.28 | 2.316 | 2.06 | 2.183 | 2.32 | 2.147
| Std Dev. |  | 0.012111 | 0.033335 | 0.028543 | 0.0089443 | 0.017889 | 0.026833 | 0.054924 | 0.028048 | 0.01169 | 0.061563 | 0.05812 | 0.10445 | 0.01169 | 0.037639 | 0.053037 | 0.030166 | 0.10008 | 0.037947 | 0.032054 | 0.020656 | 0.019212 | 0.075476 | 0.036637
| Rel.Std.Dev. |  | 0.5615% | 1.557% | 1.231% | 0.4321% | 0.8206% | 1.278% | 2.487% | 1.296% | 0.5442% | 2.87% | 2.691% | 4.237% | 0.5223% | 1.808% | 2.467% | 1.433% | 4.567% | 1.672% | 1.389% | 1.006% | 0.8827% | 3.294% | 1.701%
| PDM3 |  | -0.8951% | -1.607% | 6.581% | -4.878% | 0.1772% | -3.499% | 1.479% | -0.5887% | -1.278% | -1.431% | -0.7395% | 13.27% | 2.858% | -4.342% | -1.188% | -3.269% | 0.7133% | 4.313% | 6.028% | -5.644% | 0.01633% | 5.309% | -1.014%
| Z-Score (Absolute) |  | 0.05 | 0.26 | 2.22 | 1.26 | 0.28 | 0.84 | 0.67 | 0.05 | 0.16 | 0.21 | 0 | 4.26 | 1.09 | 1.09 | 0.14 | 0.77 | 0.44 | 1.54 | 2.06 | 1.49 | 0.23 | 1.84 | 0.08
| Analytical results for La in OREAS 931b (Certified Value 32.3 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 24 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS |  |  |  | 1
1
| 1 | 1 | 30 | 36.6 | 32.8 | 30.3 | 31.9 | 32.2 | 28.9 | 29.6 | 38.8518 | 39 | 27.9 | 34.8 | 34.52 | 25.7 | 22.2 | 30.3 | 34.03 | 37 | 30 | 32.7151 |  |  |  | 1
| 1 | 2 | 31 | 37.21 | 32.1 | 32.4 | 31.9 | 31.9 | 30.2 | 31.7 | 38.2884 | 39.03 | 27.7 | 35.7 | 36.35 | 27.7 | 22 | 30.6 | 35.69 | 37 | 31 | 32.8936 |  |  |  | 17
| 1 | 3 | 31 | 37.01 | 33.6 | 31.6 | 30.7 | 31.1 | 31.7 | 29.6 | 37.1141 | 41.27 | 28.1 | 32.8 | 34.81 | 26.5 | 25 | 31.4 | 34.53 | 38 | 35.1 | 33.6929 |  |  |  | 16
| 1 | 4 | 30 | 37.09 | 32.8 | 30.7 | 30.1 | 32.5 | 30.8 | 32.6 | 40.6715 | 42.21 | 29.6 | 35 | 37.31 | 26.6 | 25.8 | 30.4 | 35.07 | 37 | 33.4 | 33.7089 |  |  |  | 32.3429
| 1 | 5 | 32 | 37.42 | 34.3 | 30.9 | 30.2 | 31.7 | 31.1 | 31.9 | 39.102 | 41.67 | 27.8 | 37.1 | 36.22 | 29.2 | 26 | 31.7 | 35.14 | 38 | 30.6 | 32.7647 |  |  |  | 30
| 1 | 6 | 30 | 37.26 | 33.6 | 30.7 | 29.4 | 32.3 | 31.8 | 30 | 38.6415 | 43.18 | 28.7 | 35.9 | 35.89 | 26.7 | 27.1 | 29.4 | 34.77 | 38 | 36.2 | 33.5435
| Mean |  | 30.6667 | 37.0983 | 33.2 | 31.1 | 30.7 | 31.95 | 30.75 | 30.9 | 38.7782 | 41.06 | 28.3 | 35.2167 | 35.85 | 27.0667 | 24.6833 | 30.6333 | 34.8717 | 37.5 | 32.7167 | 33.2198
| Median |  | 30.5 | 37.15 | 33.2 | 30.8 | 30.45 | 32.05 | 30.95 | 30.85 | 38.7467 | 41.47 | 28 | 35.35 | 36.055 | 26.65 | 25.4 | 30.5 | 34.92 | 37.5 | 32.2 | 33.2185
| Std Dev. |  | 0.8165 | 0.28224 | 0.7823 | 0.76681 | 1.0178 | 0.50498 | 1.0821 | 1.3206 | 1.1595 | 1.7087 | 0.72938 | 1.4359 | 1.0364 | 1.2242 | 2.1113 | 0.82624 | 0.56859 | 0.54772 | 2.5725 | 0.47667
| Rel.Std.Dev. |  | 2.662% | 0.7608% | 2.356% | 2.466% | 3.315% | 1.581% | 3.519% | 4.274% | 2.99% | 4.161% | 2.577% | 4.077% | 2.891% | 4.523% | 8.554% | 2.697% | 1.631% | 1.461% | 7.863% | 1.435%
| PDM3 |  | -5.183% | 14.7% | 2.65% | -3.843% | -5.08% | -1.215% | -4.925% | -4.461% | 19.9% | 26.95% | -12.5% | 8.885% | 10.84% | -16.31% | -23.68% | -5.286% | 7.819% | 15.95% | 1.156% | 2.711%
| Z-Score (Absolute) |  | 0.53 | 1.52 | 0.28 | 0.39 | 0.52 | 0.12 | 0.5 | 0.46 | 2.05 | 2.78 | 1.28 | 0.92 | 1.12 | 1.68 | 2.43 | 0.54 | 0.81 | 1.64 | 0.12 | 0.28
| Analytical results for Li in OREAS 931b (Certified Value 17 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 16 | 18.2753 | 17.9 | 17.1 | 17.1 | 16.6 | 17.9 | 17.2 | 17 | 14 | 17.193 | 23 | 16.2 | 16.5 | 16.6 | 17.3 | 17 | 18.1 | 13.0507 | 14.2 | 16 | 18 | 16.536 | 1
| 1 | 2 | 15 | 17.8417 | 17.9 | 16.9 | 17.6 | 16.2 | 17.7 | 16.8 | 17.3 | 15 | 17.3492 | 24 | 16.2 | 16.9 | 16.5 | 16.6 | 16 | 17.9 | 13.5845 | 14.9 | 16 | 18 | 16.692 | 0
| 1 | 3 | 16 | 17.2528 | 17.8 | 16.8 | 17 | 16.4 | 17.4 | 16.5 | 17.5 | 14 | 16.1111 | 26 | 16.2 | 16.2 | 15.2 | 18.2 | 16 | 18.1 | 13.6231 | 14.8 | 16 | 18 | 17.238 | 16
| 1 | 4 | 16 | 17.8125 | 17.7 | 17 | 17.1 | 16.2 | 17.1 | 16.8 | 17.4 | 15 | 17.6765 | 26 | 16.3 | 17.2 | 15.9 | 17.3 | 17 | 17.3 | 13.3257 | 14.1 | 16 | 18 | 17.238 | 16.9919
| 1 | 5 | 16 | 17.7236 | 17.5 | 17.2 | 17.4 | 16.5 | 16.8 | 16.9 | 17.3 | 15 | 16.4861 | 29 | 16.3 | 17.4 | 15.4 | 16.1 | 17 | 17.8 | 12.9135 | 15.5 | 16 | 17 | 16.614 | 31
| 1 | 6 | 16 | 17.2106 | 17.2 | 16.3 | 17.2 | 16.4 | 17.2 | 19.6 | 17.2 | 15 | 16.3946 | 27 | 16.1 | 17.2 | 15.5 | 17.4 | 17 | 16.8 | 13.3667 | 14.7 | 16 | 18 | 17.082
| Mean |  | 15.8333 | 17.6861 | 17.6667 | 16.8833 | 17.2333 | 16.3833 | 17.35 | 17.3 | 17.2833 | 14.6667 | 16.8684 | 25.8333 | 16.2167 | 16.9 | 15.85 | 17.15 | 16.6667 | 17.6667 | 13.3107 | 14.7 | 16 | 17.8333 | 16.9
| Median |  | 16 | 17.768 | 17.75 | 16.95 | 17.15 | 16.4 | 17.3 | 16.85 | 17.3 | 15 | 16.8395 | 26 | 16.2 | 17.05 | 15.7 | 17.3 | 17 | 17.85 | 13.3462 | 14.75 | 16 | 18 | 16.887
| Std Dev. |  | 0.40825 | 0.40066 | 0.27325 | 0.31885 | 0.22509 | 0.16021 | 0.40373 | 1.1489 | 0.17224 | 0.5164 | 0.62188 | 2.137 | 0.075277 | 0.46476 | 0.58907 | 0.72319 | 0.5164 | 0.5164 | 0.28331 | 0.5099 | 0 | 0.40825 | 0.32223
| Rel.Std.Dev. |  | 2.578% | 2.265% | 1.547% | 1.889% | 1.306% | 0.9779% | 2.327% | 6.641% | 0.9966% | 3.521% | 3.687% | 8.272% | 0.4642% | 2.75% | 3.717% | 4.217% | 3.098% | 2.923% | 2.128% | 3.469% | 0% | 2.289% | 1.907%
| PDM3 |  | -6.818% | 4.086% | 3.971% | -0.6387% | 1.421% | -3.581% | 2.108% | 1.813% | 1.715% | -13.68% | -0.7265% | 52.03% | -4.562% | -0.5406% | -6.72% | 0.9307% | -1.914% | 3.971% | -21.66% | -13.49% | -5.837% | 4.952% | -0.5406%
| Z-Score (Absolute) |  | NA | 1.18 | 1.15 | 0.02 | 0.5 | 0.77 | 0.67 | 0.6 | 0.57 | NA | 0.05 | NA | 1.02 | 0 | 1.57 | 0.37 | NA | 1.15 | 5.38 | 3.3 | NA | NA | 0
| Analytical results for Lu in OREAS 931b (Certified Value 0.22 ppm) | [24 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 03; 02: 10; 03: 11; 04: 12; 05: 13; 06: 15; 07: 17; 08: 20; 09: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*OES/MS | 2
3
| 1 | 1 | 01: 0.22; 02: 0.22; 03: 0.2214; 04: 0.23; 05: 0.2; 06: 0.2; 07: 0.3; 08: 0.22; 09: 0.1254 | 1
| 1 | 2 | 01: 0.21; 02: 0.24; 03: 0.2184; 04: 0.23; 05: 0.2; 06: 0.22; 07: 0.3; 08: 0.22; 09: 0.1111 | 0
| 1 | 3 | 01: 0.21; 02: 0.22; 03: 0.2105; 04: 0.24; 05: 0.2; 06: 0.2; 07: 0.3; 08: 0.21; 09: 0.1099 | 16
| 1 | 4 | 01: 0.21; 02: 0.25; 03: 0.2272; 04: 0.25; 05: 0.2; 06: 0.2; 07: 0.3; 08: 0.22; 09: 0.1124 | 0.22246
| 1 | 5 | 01: 0.21; 02: 0.25; 03: 0.2227; 04: 0.25; 05: 0.2; 06: 0.21; 07: 0.3; 08: 0.23; 09: 0.0997 | 32
| 1 | 6 | 01: 0.21; 02: 0.23; 03: 0.2185; 04: 0.27; 05: 0.2; 06: 0.2; 07: 0.3; 08: 0.21; 09: 0.1037
| Mean |  | 01: 0.21167; 02: 0.235; 03: 0.21978; 04: 0.245; 05: 0.2; 06: 0.205; 07: 0.3; 08: 0.21833; 09: 0.11037
| Median |  | 01: 0.21; 02: 0.235; 03: 0.21995; 04: 0.245; 05: 0.2; 06: 0.2; 07: 0.3; 08: 0.22; 09: 0.1105
| Std Dev. |  | 01: 0.0040825; 02: 0.013784; 03: 0.0055833; 04: 0.015166; 05: 3.0405e-17; 06: 0.0083666; 07: 0; 08: 0.0075277; 09: 0.0088217
| Rel.Std.Dev. |  | 01: 1.929%; 02: 5.866%; 03: 2.54%; 04: 6.19%; 05: 1.52e-14%; 06: 4.081%; 07: 0%; 08: 3.448%; 09: 7.993%
| PDM3 |  | 01: -4.853%; 02: 5.635%; 03: -1.205%; 04: 10.13%; 05: -10.1%; 06: -7.85%; 07: 34.85%; 08: -1.857%; 09: -50.39%
| Z-Score (Absolute) |  | 01: 0.34; 02: 0.84; 03: 0.07; 04: 1.35; 05: NA; 06: 0.67; 07: NA; 08: 0; 09: 5.46
Results from laboratories 13 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 931b (Certified Value 0.354 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.36 | 0.37813 | 0.3873 | 0.359 | 0.35 | 0.34 | 0.35 | 0.36 | 0.34 | 0.32 | 0.38291 | 0.406 | 0.36 | 0.34 | 0.3393 | 0.32 | 0.35 | 0.38 | 0.3784 | 0.358 | 0.362 | 0.36 | 0.35801 | 1
| 1 | 2 | 0.35 | 0.36266 | 0.3751 | 0.358 | 0.35 | 0.33 | 0.35 | 0.35 | 0.34 | 0.33 | 0.3761 | 0.413 | 0.35 | 0.34 | 0.3461 | 0.33 | 0.35 | 0.37 | 0.3914 | 0.356 | 0.356 | 0.37 | 0.35967 | 0
| 1 | 3 | 0.36 | 0.37828 | 0.3834 | 0.36 | 0.35 | 0.32 | 0.34 | 0.34 | 0.34 | 0.31 | 0.36273 | 0.415 | 0.35 | 0.33 | 0.3255 | 0.32 | 0.37 | 0.39 | 0.3822 | 0.35 | 0.356 | 0.36 | 0.37035 | 16
| 1 | 4 | 0.37 | 0.36463 | 0.391 | 0.356 | 0.35 | 0.32 | 0.34 | 0.35 | 0.34 | 0.33 | 0.39021 | 0.42 | 0.35 | 0.33 | 0.3499 | 0.32 | 0.36 | 0.38 | 0.3827 | 0.351 | 0.356 | 0.37 | 0.36747 | 0.35396
| 1 | 5 | 0.36 | 0.36644 | 0.3819 | 0.353 | 0.35 | 0.32 | 0.33 | 0.35 | 0.34 | 0.33 | 0.37378 | 0.443 | 0.35 | 0.33 | 0.3441 | 0.33 | 0.36 | 0.38 | 0.3781 | 0.358 | 0.356 | 0.38 | 0.35755 | 33
| 1 | 6 | 0.35 | 0.37375 | 0.378 | 0.357 | 0.35 | 0.32 | 0.34 | 0.35 | 0.34 | 0.3 | 0.37181 | 0.451 | 0.35 | 0.33 | 0.3347 | 0.33 | 0.36 | 0.39 | 0.3839 | 0.356 | 0.362 | 0.36 | 0.36637
| Mean |  | 0.35833 | 0.37065 | 0.38278 | 0.35717 | 0.35 | 0.325 | 0.34167 | 0.35 | 0.34 | 0.32 | 0.37625 | 0.42467 | 0.35167 | 0.33333 | 0.33993 | 0.325 | 0.35833 | 0.38167 | 0.38278 | 0.35483 | 0.358 | 0.36667 | 0.36324
| Median |  | 0.36 | 0.3701 | 0.38265 | 0.3575 | 0.35 | 0.32 | 0.34 | 0.35 | 0.34 | 0.325 | 0.37494 | 0.4175 | 0.35 | 0.33 | 0.3417 | 0.325 | 0.36 | 0.38 | 0.38245 | 0.356 | 0.356 | 0.365 | 0.36302
| Std Dev. |  | 0.0075277 | 0.0069501 | 0.0058479 | 0.0024833 | 6.0809e-17 | 0.0083666 | 0.0075277 | 0.0063246 | 0 | 0.012649 | 0.0094632 | 0.018052 | 0.0040825 | 0.005164 | 0.0088412 | 0.0054772 | 0.0075277 | 0.0075277 | 0.0048371 | 0.0034881 | 0.0030984 | 0.008165 | 0.0054911
| Rel.Std.Dev. |  | 2.101% | 1.875% | 1.528% | 0.6953% | 1.737e-14% | 2.574% | 2.203% | 1.807% | 0% | 3.953% | 2.515% | 4.251% | 1.161% | 1.549% | 2.601% | 1.685% | 2.101% | 1.972% | 1.264% | 0.983% | 0.8655% | 2.227% | 1.512%
| PDM3 |  | 1.235% | 4.715% | 8.143% | 0.9059% | -1.119% | -8.182% | -3.473% | -1.119% | -3.944% | -9.594% | 6.299% | 19.98% | -0.648% | -5.827% | -3.963% | -8.182% | 1.235% | 7.828% | 8.143% | 0.2467% | 1.141% | 3.59% | 2.621%
| Z-Score (Absolute) |  | 0.05 | 0.59 | 1.11 | 0 | 0.31 | 1.4 | 0.67 | 0.31 | 0.75 | 1.62 | 0.83 | 2.94 | 0.24 | 1.04 | 0.75 | 1.4 | 0.05 | 1.07 | 1.11 | 0.1 | 0.03 | 0.41 | 0.26
| Analytical results for Mn in OREAS 931b (Certified Value 0.016 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.0241 | 0.016868 | 0.0173 | 0.0159 | 0.0161 | 0.016 | 0.0167 | 0.0163 | 0.0161 | 0.02 | 0.017156 | 0.0181 | 0.017 | 0.0162 | 0.0164 | 0.0149 | 0.0156 | 0.0166 | 0.01557 | 0.0164 | 0.015 | 0.0162 | 0.014858 | 1
| 1 | 2 | 0.0244 | 0.016913 | 0.0174 | 0.0158 | 0.0164 | 0.0156 | 0.0165 | 0.016 | 0.0163 | 0.02 | 0.015566 | 0.0183 | 0.0171 | 0.0166 | 0.0169 | 0.0152 | 0.0153 | 0.0168 | 0.01598 | 0.0163 | 0.015 | 0.0162 | 0.014861 | 0
| 1 | 3 | 0.0244 | 0.016911 | 0.0173 | 0.0156 | 0.0161 | 0.0163 | 0.0164 | 0.0157 | 0.0164 | 0.02 | 0.014803 | 0.0186 | 0.0169 | 0.0163 | 0.0161 | 0.0152 | 0.0149 | 0.0169 | 0.015984 | 0.0161 | 0.015 | 0.0165 | 0.015398 | 16
| 1 | 4 | 0.0245 | 0.017066 | 0.0174 | 0.0161 | 0.0162 | 0.0163 | 0.0162 | 0.016 | 0.0161 | 0.02 | 0.016091 | 0.0186 | 0.0169 | 0.0162 | 0.0173 | 0.015 | 0.0202 | 0.0167 | 0.015961 | 0.0161 | 0.015 | 0.0162 | 0.015076 | 0.016092
| 1 | 5 | 0.0243 | 0.016956 | 0.017 | 0.0155 | 0.0164 | 0.0161 | 0.0156 | 0.016 | 0.0162 | 0.02 | 0.015432 | 0.0196 | 0.017 | 0.016 | 0.0167 | 0.0154 | 0.0193 | 0.0167 | 0.015705 | 0.0164 | 0.015 | 0.0155 | 0.014647 | 34
| 1 | 6 | 0.0243 | 0.01712 | 0.0169 | 0.0155 | 0.0165 | 0.0162 | 0.0159 | 0.0159 | 0.0162 | 0.01 | 0.015297 | 0.02 | 0.0168 | 0.0161 | 0.0164 | 0.0156 | 0.0206 | 0.0168 | 0.015772 | 0.0163 | 0.015 | 0.0159 | 0.014862
| Mean |  | 0.024333 | 0.016972 | 0.017217 | 0.015733 | 0.016283 | 0.016083 | 0.016217 | 0.015983 | 0.016217 | 0.018333 | 0.015724 | 0.018867 | 0.01695 | 0.016233 | 0.016633 | 0.015217 | 0.01765 | 0.01675 | 0.015829 | 0.016267 | 0.015 | 0.016083 | 0.014951
| Median |  | 0.02435 | 0.016935 | 0.0173 | 0.0157 | 0.0163 | 0.01615 | 0.0163 | 0.016 | 0.0162 | 0.02 | 0.015499 | 0.0186 | 0.01695 | 0.0162 | 0.01655 | 0.0152 | 0.01745 | 0.01675 | 0.015867 | 0.0163 | 0.015 | 0.0162 | 0.014862
| Std Dev. |  | 0.00013663 | 9.8853e-05 | 0.0002137 | 0.00024221 | 0.00017224 | 0.00026394 | 0.00040702 | 0.00019408 | 0.0001169 | 0.0040825 | 0.00081586 | 0.00075807 | 0.00010488 | 0.00020656 | 0.0004274 | 0.00025626 | 0.0026539 | 0.00010488 | 0.0001731 | 0.00013663 | 0 | 0.00034303 | 0.00025794
| Rel.Std.Dev. |  | 0.5615% | 0.5824% | 1.241% | 1.539% | 1.058% | 1.641% | 2.51% | 1.214% | 0.7209% | 22.27% | 5.189% | 4.018% | 0.6188% | 1.272% | 2.57% | 1.684% | 15.04% | 0.6262% | 1.094% | 0.8399% | 0% | 2.133% | 1.725%
| PDM3 |  | 51.21% | 5.471% | 6.989% | -2.229% | 1.189% | -0.05357% | 0.775% | -0.675% | 0.775% | 13.93% | -2.285% | 17.24% | 5.332% | 0.8786% | 3.364% | -5.439% | 9.682% | 4.089% | -1.635% | 1.086% | -6.786% | -0.05357% | -7.093%
| Z-Score (Absolute) |  | 10.73 | 0.98 | 1.3 | 0.66 | 0.07 | 0.2 | 0.02 | 0.33 | 0.02 | 2.78 | 0.67 | 3.49 | 0.95 | 0 | 0.53 | 1.35 | 1.88 | 0.68 | 0.54 | 0.04 | 0.99 | 0.2 | 1.7
| Analytical results for Mo in OREAS 931b (Certified Value 3.38 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 3 | 3.2194 | 3.6 | 3.3 | 3.17 | 3.21 | 3.17 | 3.66 | 3.25 | 3.41 | < 5 | 3.18 | 3.59 | 3.53 | 3.3 | 2.8 | 2 | 3.5 | 3.3535 | 3.41 | 4 | 3 | 1.8368 | 1
| 1 | 2 | 3 | 3.2301 | 3.4 | 3.4 | 3.16 | 3.3 | 3.2 | 3.67 | 3.4 | 3.56 | < 5 | 3.23 | 3.54 | 3.55 | 3.2 | 3.1 | 1 | 3.39 | 3.4907 | 3.5 | 4 | 3 | 1.7682 | 8
| 1 | 3 | 3 | 3.2821 | 3.5 | 3.3 | 3.41 | 3.53 | 2.97 | 3.52 | 3.38 | 3.5 | < 5 | 3.4 | 3.63 | 3.46 | 3.1 | 3.2 | 3 | 3.55 | 3.4425 | 3.42 | 6 | 3 | 1.7997 | 16
| 1 | 4 | 3 | 3.2221 | 3.5 | 3.4 | 3.04 | 3.57 | 3.07 | 3.66 | 3.26 | 3.52 | < 5 | 3.53 | 3.5 | 3.65 | 3.4 | 3.2 | 1 | 3.39 | 3.4447 | 3.33 | 4 | 3 | 1.8591 | 3.3832
| 1 | 5 | 3 | 3.3121 | 3.3 | 3.4 | 3.15 | 3.5 | 2.95 | 3.55 | 3.36 | 3.53 | < 5 | 3.44 | 3.76 | 3.6 | 3.3 | 3.5 | 1 | 3.55 | 3.4085 | 3.44 | 4 | 3 | 1.6627 | 35
| 1 | 6 | 3 | 3.1728 | 3.3 | 3.3 | 3.25 | 3.47 | 2.92 | 3.7 | 3.22 | 3.49 | < 5 | 3.53 | 3.67 | 3.54 | 3.3 | 3.5 | < 1 | 3.35 | 3.3889 | 3.43 | 5 | 3 | 1.7406
| Mean |  | 3 | 3.2397 | 3.4333 | 3.35 | 3.1967 | 3.43 | 3.0467 | 3.6267 | 3.3117 | 3.5017 |  | 3.385 | 3.615 | 3.555 | 3.2667 | 3.2167 | 1.6 | 3.455 | 3.4215 | 3.4217 | 4.5 | 3 | 1.7778
| Median |  | 3 | 3.2261 | 3.45 | 3.35 | 3.165 | 3.485 | 3.02 | 3.66 | 3.31 | 3.51 |  | 3.42 | 3.61 | 3.545 | 3.3 | 3.2 | 1 | 3.445 | 3.4255 | 3.425 | 4 | 3 | 1.7839
| Std Dev. |  | 0 | 0.049651 | 0.12111 | 0.054772 | 0.1242 | 0.14241 | 0.11877 | 0.073121 | 0.07705 | 0.051153 |  | 0.14923 | 0.093541 | 0.06473 | 0.10328 | 0.26394 | 0.89443 | 0.088938 | 0.04822 | 0.054924 | 0.83666 | 0 | 0.071126
| Rel.Std.Dev. |  | 0% | 1.533% | 3.527% | 1.635% | 3.885% | 4.152% | 3.898% | 2.016% | 2.327% | 1.461% |  | 4.409% | 2.588% | 1.821% | 3.162% | 8.206% | 55.9% | 2.574% | 1.409% | 1.605% | 18.59% | 0% | 4.001%
| PDM3 |  | -11.33% | -4.239% | 1.483% | -0.9799% | -5.512% | 1.385% | -9.946% | 7.198% | -2.113% | 3.503% |  | 0.05468% | 6.853% | 5.08% | -3.443% | -4.921% | -52.71% | 2.124% | 1.133% | 1.138% | 33.01% | -11.33% | -47.45%
| Z-Score (Absolute) |  | NA | 0.83 | 0.28 | 0.2 | 1.07 | 0.26 | 1.93 | 1.38 | 0.42 | 0.66 | 5.04 | 0 | 1.31 | 0.97 | 0.67 | 0.96 | NA | 0.4 | 0.21 | 0.21 | NA | NA | 9.16
Results from laboratories 1, 17, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 931b (Certified Value 0.076 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.07 | 0.072623 | 0.0858 | 0.078 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.093048 | 0.114 | 0.081 | 0.08 | 0.0795 | 0.061 | 0.08 | 0.06 | 0.07581 | 0.071 | 0.074 | 0.08 | 0.0597 | 1
| 1 | 2 | 0.07 | 0.072291 | 0.0829 | 0.08 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.091459 | 0.124 | 0.082 | 0.08 | 0.08 | 0.063 | 0.08 | 0.05 | 0.07661 | 0.072 | 0.074 | 0.08 | 0.0597 | 19
| 1 | 3 | 0.08 | 0.074913 | 0.0838 | 0.078 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.088647 | 0.145 | 0.081 | 0.08 | 0.0817 | 0.061 | 0.08 | 0.07 | 0.07555 | 0.071 | 0.074 | 0.08 | 0.0625 | 16
| 1 | 4 | 0.07 | 0.072456 | 0.0851 | 0.078 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.095899 | 0.161 | 0.081 | 0.08 | 0.0779 | 0.061 | 0.07 | 0.06 | 0.07676 | 0.071 | 0.074 | 0.08 | 0.0618 | 0.076484
| 1 | 5 | 0.08 | 0.073171 | 0.0824 | 0.077 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.089773 | 0.199 | 0.082 | 0.08 | 0.0796 | 0.06 | 0.08 | 0.07 | 0.07326 | 0.071 | 0.074 | 0.07 | 0.0598 | 36
| 1 | 6 | 0.07 | 0.074992 | 0.0822 | 0.076 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.06 | 0.090357 | 0.157 | 0.08 | 0.08 | 0.0757 | 0.064 | 0.07 | 0.07 | 0.07622 | 0.071 | 0.074 | 0.09 | 0.061
| Mean |  | 0.073333 | 0.073408 | 0.0837 | 0.077833 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.068333 | 0.09153 | 0.15 | 0.081167 | 0.08 | 0.079067 | 0.061667 | 0.076667 | 0.063333 | 0.075702 | 0.071167 | 0.074 | 0.08 | 0.06075
| Median |  | 0.07 | 0.072897 | 0.08335 | 0.078 | 0.07 | 0.07 | 0.08 | 0.08 | 0.08 | 0.07 | 0.090908 | 0.151 | 0.081 | 0.08 | 0.07955 | 0.061 | 0.08 | 0.065 | 0.076015 | 0.071 | 0.074 | 0.08 | 0.0604
| Std Dev. |  | 0.005164 | 0.0012329 | 0.0014805 | 0.0013292 | 0 | 0 | 0 | 0 | 0 | 0.0040825 | 0.0026143 | 0.030226 | 0.00075277 | 0 | 0.0020481 | 0.0015055 | 0.005164 | 0.008165 | 0.0012814 | 0.00040825 | 0 | 0.0063246 | 0.0012112
| Rel.Std.Dev. |  | 7.042% | 1.679% | 1.769% | 1.708% | 0% | 0% | 0% | 0% | 0% | 5.974% | 2.856% | 20.15% | 0.9274% | 0% | 2.59% | 2.441% | 6.736% | 12.89% | 1.693% | 0.5737% | 0% | 7.906% | 1.994%
| PDM3 |  | -4.12% | -4.022% | 9.434% | 1.764% | -8.478% | -8.478% | 4.597% | 4.597% | 4.597% | -10.66% | 19.67% | 96.12% | 6.122% | 4.597% | 3.377% | -19.37% | 0.2386% | -17.19% | -1.023% | -6.952% | -3.248% | 4.597% | -20.57%
| Z-Score (Absolute) |  | 0.67 | 0.66 | 1.42 | 0.24 | 1.35 | 1.35 | 0.67 | 0.67 | 0.67 | 1.69 | 3.01 | 14.83 | 0.91 | 0.67 | 0.49 | 3.03 | 0 | 2.7 | 0.2 | 1.11 | 0.5 | 0.67 | 3.22
| Analytical results for Nb in OREAS 931b (Certified Value 2.72 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 2
2
| 1 | 1 | < 5 | 3.6471 | 3.02 | 2 | 2.7 | 3 | 2.8 | 2.8 | 2.8 | 4 | 1.8 | 2.38 | 2.3 | 2.81 | 1.8 | 8.7 | 3.1 | 2.7896 | 2.2 | 23 | 1.8368 |  |  | 1
| 1 | 2 | < 5 | 3.6564 | 3.41 | 2 | 2.5 | 3.4 | 2.9 | 2.8 | 2.9 | 3.9 | 1.8 | 2.39 | 2.3 | 2.72 | 2.9 | 9 | 3.1 | 2.8281 | 2.3 | 24 | 1.7682 |  |  | 10
| 1 | 3 | < 5 | 3.5014 | 3.94 | 2.2 | 1.9 | 3.5 | 3 | 3.1 | 2.8 | 3.7 | 1.9 | 2.29 | 2.2 | 2.89 | 1.8 | 9.4 | 3.1 | 2.8531 | 2.2 | 24 | 1.7997 |  |  | 16
| 1 | 4 | < 5 | 3.2708 | 3.73 | 2.1 | 2.7 | 3.3 | 2.8 | 2.7 | 2.8 | 4 | 2 | 2.18 | 2.2 | 2.89 | 1.7 | 8.2 | 3 | 2.8128 | 2.2 | 23 | 1.8591 |  |  | 2.7246
| 1 | 5 | < 5 | 3.2492 | 3.74 | 2.4 | 2.6 | 3.4 | 2.5 | 2.8 | 3.1 | 3.8 | 2.1 | 2.29 | 2.3 | 2.72 | 2 | 7.6 | 3.1 | 2.563 | 2.1 | 23 | 1.6627 |  |  | 37
| 1 | 6 | < 5 | 3.4274 | 4.47 | 2.1 | 2.6 | 3.4 | 2.7 | 2.8 | 2.9 | 4.7 | 2 | 2.28 | 2.2 | 3.13 | 2.5 | 8 | 3.1 | 2.9448 | 2.3 | 24 | 1.7406
| Mean |  |  | 3.4587 | 3.7183 | 2.1333 | 2.5 | 3.3333 | 2.7833 | 2.8333 | 2.8833 | 4.0167 | 1.9333 | 2.3017 | 2.25 | 2.86 | 2.1167 | 8.4833 | 3.0833 | 2.7986 | 2.2167 | 23.5 | 1.7778
| Median |  |  | 3.4644 | 3.735 | 2.1 | 2.6 | 3.4 | 2.8 | 2.8 | 2.85 | 3.95 | 1.95 | 2.29 | 2.25 | 2.85 | 1.9 | 8.45 | 3.1 | 2.8205 | 2.2 | 23.5 | 1.7839
| Std Dev. |  |  | 0.17698 | 0.48938 | 0.15055 | 0.30332 | 0.17512 | 0.17224 | 0.13663 | 0.1169 | 0.35449 | 0.12111 | 0.07679 | 0.054772 | 0.15258 | 0.47924 | 0.67057 | 0.040825 | 0.12727 | 0.075277 | 0.54772 | 0.071126
| Rel.Std.Dev. |  |  | 5.117% | 13.16% | 7.057% | 12.13% | 5.254% | 6.188% | 4.822% | 4.054% | 8.826% | 6.264% | 3.336% | 2.434% | 5.335% | 22.64% | 7.905% | 1.324% | 4.548% | 3.396% | 2.331% | 4.001%
| PDM3 |  |  | 26.95% | 36.47% | -21.7% | -8.242% | 22.34% | 2.157% | 3.992% | 5.827% | 47.42% | -29.04% | -15.52% | -17.42% | 4.971% | -22.31% | 211.4% | 13.17% | 2.716% | -18.64% | 762.5% | -34.75%
| Z-Score (Absolute) |  | 0.37 | 0.81 | 1.13 | 0.82 | 0.37 | 0.66 | 0.02 | 0.04 | 0.1 | 1.5 | 1.06 | 0.61 | 0.67 | 0.08 | 0.84 | 6.99 | 0.35 | 0 | 0.72 | 25.45 | 1.25
| Analytical results for Nd in OREAS 931b (Certified Value 29.9 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  | 1
| No | No | 01: 03; 02: 11; 03: 12; 04: 13; 05: 15; 06: 17; 07: 22; 08: 24 |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES; 08: 4A*OES/MS |  |  | 1
2
| 1 | 1 | 01: 31.75; 02: 27.5715; 03: 32.1; 04: 26.7; 05: 29.78; 06: 29.7; 07: 32; 08: 27.6875 |  |  | 1
| 1 | 2 | 01: 32.02; 02: 27.2016; 03: 32.6; 04: 26.7; 05: 30.97; 06: 30.4; 07: 32; 08: 27.8142 |  |  | 5
| 1 | 3 | 01: 31.64; 02: 26.2553; 03: 34; 04: 27.4; 05: 30.33; 06: 30.9; 07: 32; 08: 27.3365 |  |  | 16
| 1 | 4 | 01: 31.56; 02: 28.5645; 03: 34; 04: 26.6; 05: 32.49; 06: 30.8; 07: 32; 08: 28.7173 |  |  | 29.8611
| 1 | 5 | 01: 31.93; 02: 27.7107; 03: 33.2; 04: 26.5; 05: 31.35; 06: 32.1; 07: 32; 08: 26.7999 |  |  | 38
| 1 | 6 | 01: 31.93; 02: 27.5023; 03: 35; 04: 26; 05: 30.7; 06: 32; 07: 33; 08: 27.8563
| Mean |  | 01: 31.805; 02: 27.4676; 03: 33.4833; 04: 26.65; 05: 30.9367; 06: 30.9833; 07: 32.1667; 08: 27.702
| Median |  | 01: 31.84; 02: 27.5369; 03: 33.6; 04: 26.65; 05: 30.835; 06: 30.85; 07: 32; 08: 27.7508
| Std Dev. |  | 01: 0.18317; 02: 0.7503; 03: 1.0591; 04: 0.45056; 05: 0.93239; 06: 0.92826; 07: 0.40825; 08: 0.63479
| Rel.Std.Dev. |  | 01: 0.5759%; 02: 2.732%; 03: 3.163%; 04: 1.691%; 05: 3.014%; 06: 2.996%; 07: 1.269%; 08: 2.292%
| PDM3 |  | 01: 6.51%; 02: -8.015%; 03: 12.13%; 04: -10.75%; 05: 3.602%; 06: 3.758%; 07: 7.721%; 08: -7.231%
| Z-Score (Absolute) |  | 01: 0.23; 02: 0.92; 03: 0.67; 04: 1.14; 05: 0; 06: 0.01; 07: NA; 08: 0.86
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ni in OREAS 931b (Certified Value 12.6 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 14 | 12.893 | 12.6 | 12 | 12.2 | 11.3 | 12.1 | 13.3 | 12.8 | 12.4 | 12.4041 | 17 | 12.8 | 12.2 | 12.3 | 12.6 | 12 | 13.3 | 12.1648 | 11 | 14 | 11 | 9.349 | 1
| 1 | 2 | 13 | 12.8492 | 12.4 | 12 | 12.3 | 11.4 | 11.9 | 12.4 | 13.4 | 12.7 | 12.6637 | 21 | 10.8 | 13.1 | 12.7 | 12.9 | 12 | 13.9 | 12.8103 | 10 | 14 | 12 | 10.205 | 12
| 1 | 3 | 15 | 12.7257 | 12.9 | 12 | 12.4 | 13 | 11.4 | 12.3 | 13.3 | 12.4 | 12.3364 | 17 | 10.6 | 12.4 | 12.3 | 12.4 | 14 | 13.7 | 12.2913 | 10 | 13 | 11 | 8.9851 | 16
| 1 | 4 | 13 | 12.8453 | 13.1 | 12 | 11.6 | 13.7 | 11.3 | 13.1 | 13 | 12.6 | 14.5587 | 17 | 10.5 | 12.2 | 11.9 | 13 | 15 | 13 | 12.5859 | 10 | 14 | 11 | 9.1696 | 12.5914
| 1 | 5 | 14 | 12.939 | 12.6 | 12 | 12 | 13.1 | 10.9 | 12.8 | 13.6 | 12.7 | 14.0113 | 18 | 10.8 | 12.2 | 12.5 | 13 | 14 | 13.5 | 12.1549 | 11 | 14 | 14 | 8.4417 | 39
| 1 | 6 | 13 | 12.7409 | 12.2 | 12 | 11.8 | 12.6 | 11 | 12.7 | 14 | 12.1 | 11.7338 | 21 | 10.8 | 12.2 | 12.6 | 12.5 | 14 | 13.8 | 12.2329 | 10 | 15 | 11 | 8.534
| Mean |  | 13.6667 | 12.8322 | 12.6333 | 12 | 12.05 | 12.5167 | 11.4333 | 12.7667 | 13.35 | 12.4833 | 12.9513 | 18.5 | 11.05 | 12.3833 | 12.3833 | 12.7333 | 13.5 | 13.5333 | 12.3734 | 10.3333 | 14 | 11.6667 | 9.1141
| Median |  | 13.5 | 12.8472 | 12.6 | 12 | 12.1 | 12.8 | 11.35 | 12.75 | 13.35 | 12.5 | 12.5339 | 17.5 | 10.8 | 12.2 | 12.4 | 12.75 | 14 | 13.6 | 12.2621 | 10 | 14 | 11 | 9.0773
| Std Dev. |  | 0.8165 | 0.083931 | 0.3266 | 0 | 0.30822 | 0.9704 | 0.48028 | 0.38816 | 0.42778 | 0.23166 | 1.0909 | 1.9748 | 0.8666 | 0.36009 | 0.28577 | 0.26583 | 1.2247 | 0.33862 | 0.26598 | 0.5164 | 0.63246 | 1.2111 | 0.64087
| Rel.Std.Dev. |  | 5.974% | 0.6541% | 2.585% | 0% | 2.558% | 7.753% | 4.201% | 3.04% | 3.204% | 1.856% | 8.423% | 10.67% | 7.843% | 2.908% | 2.308% | 2.088% | 9.072% | 2.502% | 2.15% | 4.997% | 4.518% | 10.38% | 7.032%
| PDM3 |  | 8.539% | 1.912% | 0.3327% | -4.697% | -4.3% | -0.5939% | -9.198% | 1.392% | 6.024% | -0.8586% | 2.858% | 46.93% | -12.24% | -1.653% | -1.653% | 1.127% | 7.216% | 7.48% | -1.732% | -17.93% | 11.19% | -7.344% | -27.62%
| Z-Score (Absolute) |  | NA | 0.75 | 0.3 | NA | 1.01 | 0.04 | 2.4 | 0.6 | 1.91 | 0.04 | 1.02 | NA | 3.27 | 0.26 | 0.26 | 0.53 | NA | 2.33 | 0.29 | NA | NA | NA | 7.63
Results from laboratories 1, 4, 12, 17, 20, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 931b (Certified Value 0.027 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.0254 | 0.026721 | 0.0289 | 0.028 | 0.026 | 0.027 | 0.027 | 0.028 | 0.027 | 0.0251 | 0.029987 | 0.027 | 0.027 | 0.0382 | 0.0287 | 0.028 | 0.04 | 0.0284 | 0.0898 | 0.027 | 0.087 | 0.03 | 0.025296 | 1
| 1 | 2 | 0.0242 | 0.024685 | 0.0286 | 0.028 | 0.027 | 0.026 | 0.027 | 0.027 | 0.027 | 0.0264 | 0.029087 | 0.027 | 0.0273 | 0.0409 | 0.0288 | 0.028 | 0.04 | 0.0274 | 0.1003 | 0.027 | 0.087 | 0.03 | 0.024996 | 0
| 1 | 3 | 0.0243 | 0.026843 | 0.0292 | 0.028 | 0.027 | 0.026 | 0.027 | 0.027 | 0.027 | 0.0245 | 0.027571 | 0.028 | 0.027 | 0.04 | 0.0292 | 0.027 | 0.04 | 0.029 | 0.09269 | 0.027 | 0.083 | 0.03 | 0.026094 | 16
| 1 | 4 | 0.0237 | 0.025388 | 0.0295 | 0.028 | 0.026 | 0.026 | 0.027 | 0.027 | 0.026 | 0.0255 | 0.028623 | 0.028 | 0.0271 | 0.0386 | 0.0282 | 0.025 | 0.041 | 0.0286 | 0.09286 | 0.027 | 0.083 | 0.03 | 0.025593 | 0.027157
| 1 | 5 | 0.024 | 0.025455 | 0.0289 | 0.028 | 0.027 | 0.026 | 0.026 | 0.027 | 0.026 | 0.026 | 0.027652 | 0.03 | 0.0273 | 0.0374 | 0.029 | 0.025 | 0.04 | 0.0289 | 0.09277 | 0.027 | 0.087 | 0.03 | 0.024696 | 40
| 1 | 6 | 0.0243 | 0.026752 | 0.0285 | 0.028 | 0.027 | 0.026 | 0.026 | 0.027 | 0.026 | 0.0249 | 0.027587 | 0.03 | 0.0272 | 0.0398 | 0.0278 | 0.027 | 0.04 | 0.0291 | 0.09471 | 0.027 | 0.083 | 0.03 | 0.025894
| Mean |  | 0.024317 | 0.025974 | 0.028933 | 0.028 | 0.026667 | 0.026167 | 0.026667 | 0.027167 | 0.0265 | 0.0254 | 0.028418 | 0.028333 | 0.02715 | 0.03915 | 0.028617 | 0.026667 | 0.040167 | 0.028567 | 0.093855 | 0.027 | 0.085 | 0.03 | 0.025428
| Median |  | 0.02425 | 0.026088 | 0.0289 | 0.028 | 0.027 | 0.026 | 0.027 | 0.027 | 0.0265 | 0.0253 | 0.028137 | 0.028 | 0.02715 | 0.0392 | 0.02875 | 0.027 | 0.04 | 0.02875 | 0.092815 | 0.027 | 0.085 | 0.03 | 0.025445
| Std Dev. |  | 0.00057764 | 0.00091572 | 0.00037238 | 0 | 0.0005164 | 0.00040825 | 0.0005164 | 0.00040825 | 0.00054772 | 0.00070993 | 0.00099451 | 0.0013663 | 0.00013784 | 0.0013019 | 0.00052313 | 0.0013663 | 0.00040825 | 0.00062823 | 0.0035281 | 0 | 0.0021909 | 0 | 0.00053415
| Rel.Std.Dev. |  | 2.375% | 3.526% | 1.287% | 0% | 1.936% | 1.56% | 1.936% | 1.503% | 2.067% | 2.795% | 3.5% | 4.822% | 0.5077% | 3.325% | 1.828% | 5.123% | 1.016% | 2.199% | 3.759% | 0% | 2.578% | 0% | 2.101%
| PDM3 |  | -10.46% | -4.357% | 6.54% | 3.103% | -1.806% | -3.647% | -1.806% | 0.03492% | -2.42% | -6.47% | 4.642% | 4.331% | -0.02645% | 44.16% | 5.374% | -1.806% | 47.9% | 5.19% | 245.6% | -0.5788% | 213% | 10.47% | -6.367%
| Z-Score (Absolute) |  | 1.54 | 0.64 | 0.95 | 0.45 | 0.27 | 0.54 | 0.27 | 0 | 0.36 | 0.95 | 0.67 | 0.63 | 0.01 | 6.46 | 0.78 | 0.27 | 7.01 | 0.75 | 35.94 | 0.09 | 31.22 | 1.53 | 0.94
| Analytical results for Pb in OREAS 931b (Certified Value 121 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 118 | 123.3999 | 123.1 | 137 | 120.5 | 120 | 127 | 122.5 | 115.5 | 113.4 | 123.2278 | 111 | 115.18 | 119 | 118 | 124.4 | 120 | 124.7 | 102.7899 | 123 | 133 | 122 | 104.8011 | 1
| 1 | 2 | 119 | 126.5611 | 124.7 | 135 | 125 | 123 | 125.5 | 125 | 117 | 118.9 | 124.1762 | 110 | 116.59 | 118 | 122.1 | 123.5 | 100 | 121.6 | 104.4799 | 130 | 123 | 121 | 107.0693 | 13
| 1 | 3 | 116 | 126.0478 | 122.1 | 135 | 125 | 115 | 120 | 116.5 | 118.5 | 109 | 122.7433 | 121 | 115.49 | 114 | 117.7 | 124.2 | 113 | 127.1 | 104.3448 | 124 | 128 | 124 | 108.9506 | 16
| 1 | 4 | 119 | 128.6065 | 123.8 | 135 | 123 | 117 | 122 | 124.5 | 114.5 | 116.1 | 135.0043 | 119 | 115.04 | 116 | 124.2 | 122.5 | 116 | 125.9 | 102.9372 | 126 | 134 | 122 | 108.069 | 121.0905
| 1 | 5 | 117 | 127.0444 | 124.8 | 138 | 123.5 | 115 | 117 | 124 | 115 | 117.9 | 127.3099 | 116 | 116.08 | 120 | 120.5 | 121.5 | 107 | 125.7 | 99.7651 | 127 | 130 | 102 | 106.1056 | 41
| 1 | 6 | 119 | 123.2586 | 121.7 | 135 | 126 | 115 | 119 | 122.5 | 117.5 | 113.7 | 126.0366 | 122 | 114.96 | 116 | 119.4 | 125.5 | 113 | 127.3 | 100.2048 | 122 | 131 | 130 | 107.2048
| Mean |  | 118 | 125.8197 | 123.3667 | 135.8333 | 123.8333 | 117.5 | 121.75 | 122.5 | 116.3333 | 114.8333 | 126.4163 | 116.5 | 115.5567 | 117.1667 | 120.3167 | 123.6 | 111.5 | 125.3833 | 102.4203 | 125.3333 | 129.8333 | 120.1667 | 107.0334
| Median |  | 118.5 | 126.3045 | 123.45 | 135 | 124.25 | 116 | 121 | 123.25 | 116.25 | 114.9 | 125.1064 | 117.5 | 115.335 | 117 | 119.95 | 123.85 | 113 | 125.8 | 102.8636 | 125 | 130.5 | 122 | 107.137
| Std Dev. |  | 1.2649 | 2.111 | 1.3018 | 1.3292 | 1.9664 | 3.3317 | 3.8698 | 3.1145 | 1.5706 | 3.6031 | 4.5469 | 5.0892 | 0.64976 | 2.2286 | 2.5039 | 1.4311 | 7.064 | 2.0865 | 2.0153 | 2.9439 | 3.9707 | 9.4745 | 1.4576
| Rel.Std.Dev. |  | 1.072% | 1.678% | 1.055% | 0.9785% | 1.588% | 2.835% | 3.178% | 2.542% | 1.35% | 3.138% | 3.597% | 4.368% | 0.5623% | 1.902% | 2.081% | 1.158% | 6.335% | 1.664% | 1.968% | 2.349% | 3.058% | 7.884% | 1.362%
| PDM3 |  | -2.552% | 3.906% | 1.88% | 12.18% | 2.265% | -2.965% | 0.5446% | 1.164% | -3.929% | -5.167% | 4.398% | -3.791% | -4.57% | -3.24% | -0.6391% | 2.072% | -7.92% | 3.545% | -15.42% | 3.504% | 7.22% | -0.763% | -11.61%
| Z-Score (Absolute) |  | 0.39 | 0.93 | 0.52 | 2.63 | 0.6 | 0.48 | 0.24 | 0.37 | 0.67 | 0.93 | 1.03 | 0.65 | 0.81 | 0.53 | 0 | 0.56 | 1.49 | 0.86 | 3.03 | 0.85 | 1.61 | 0.03 | 2.25
| Analytical results for Pr in OREAS 931b (Certified Value 8.12 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 03; 02: 11; 03: 12; 04: 13; 05: 15; 06: 17; 07: 22; 08: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES; 08: 4A*OES/MS | 2
3
| 1 | 1 | 01: 8.47; 02: 7.5783; 03: 9.2; 04: 7.3; 05: 8.37; 06: 7.4; 07: 6; 08: 7.4345 | 1
| 1 | 2 | 01: 8.55; 02: 7.47; 03: 9.23; 04: 7.3; 05: 8.63; 06: 7.6; 07: 5; 08: 7.4961 | 14
| 1 | 3 | 01: 8.45; 02: 7.2163; 03: 9.49; 04: 7.6; 05: 8.33; 06: 7.8; 07: 5; 08: 7.3951 | 16
| 1 | 4 | 01: 8.53; 02: 7.8289; 03: 9.74; 04: 7.3; 05: 8.87; 06: 8.2; 07: 6; 08: 7.7476 | 8.1238
| 1 | 5 | 01: 8.53; 02: 7.5694; 03: 9.43; 04: 7.2; 05: 8.72; 06: 8.5; 07: 5; 08: 7.213 | 42
| 1 | 6 | 01: 8.57; 02: 7.4853; 03: 9.96; 04: 7.3; 05: 8.5; 06: 8.6; 07: 5; 08: 7.4147
| Mean |  | 01: 8.5167; 02: 7.5247; 03: 9.5083; 04: 7.3333; 05: 8.57; 06: 8.0167; 07: 5.3333; 08: 7.4502
| Median |  | 01: 8.53; 02: 7.5274; 03: 9.46; 04: 7.3; 05: 8.565; 06: 8; 07: 5; 08: 7.4246
| Std Dev. |  | 01: 0.046762; 02: 0.19853; 03: 0.29553; 04: 0.13663; 05: 0.20909; 06: 0.4916; 07: 0.5164; 08: 0.17399
| Rel.Std.Dev. |  | 01: 0.5491%; 02: 2.638%; 03: 3.108%; 04: 1.863%; 05: 2.44%; 06: 6.132%; 07: 9.682%; 08: 2.335%
| PDM3 |  | 01: 4.836%; 02: -7.375%; 03: 17.04%; 04: -9.73%; 05: 5.493%; 06: -1.319%; 07: -34.35%; 08: -8.292%
| Z-Score (Absolute) |  | 01: 0.85; 02: 0.28; 03: 1.98; 04: 0.5; 05: 0.91; 06: 0.28; 07: 2.78; 08: 0.37
| Analytical results for Rb in OREAS 931b (Certified Value 127 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  | 0
0
| 1 | 1 | 129.7427 | 127.87 | 121 | 124 | 120.5 | 132 | 129 | 130 | 119.9 | 119.4367 | 114 | 131.2 | 128 | 118.62 | 125.1 | 122 | 126.8 | 137.1675 | 127 | 129.6149 |  |  |  | 1
| 1 | 2 | 130.7101 | 127.13 | 121 | 126.5 | 119 | 132.5 | 131 | 131.5 | 126.3 | 116.1797 | 111 | 126.5 | 129 | 122.04 | 127.1 | 124 | 132.8 | 135.4477 | 131 | 129.2167 |  |  |  | 15
| 1 | 3 | 130.2223 | 127.05 | 122 | 128.5 | 123.5 | 125 | 123.5 | 133.5 | 117.3 | 110.4829 | 114 | 129.1 | 122 | 117.56 | 127.7 | 127 | 130.2 | 138.9844 | 131 | 130.1039 |  |  |  | 16
| 1 | 4 | 129.584 | 126.67 | 123 | 122.5 | 123 | 125.5 | 129 | 132.5 | 124.9 | 119.2269 | 120 | 127 | 130 | 124.59 | 125.6 | 122 | 124.8 | 138.582 | 130 | 130.6768 |  |  |  | 126.6675
| 1 | 5 | 130.9 | 126.8 | 121 | 126.5 | 123.5 | 122 | 128.5 | 134.5 | 124.2 | 115.3779 | 117 | 127.4 | 129 | 121.59 | 123.6 | 133 | 129.4 | 141.9735 | 134 | 122.019 |  |  |  | 43
| 1 | 6 | 129.6728 | 126.07 | 118 | 123.5 | 123 | 124.5 | 130.5 | 131.5 | 120.7 | 115.0284 | 117 | 127.2 | 126 | 119.29 | 126.5 | 134 | 130.1 | 141.9725 | 129 | 124.0039
| Mean |  | 130.1386 | 126.9317 | 121 | 125.25 | 122.0833 | 126.9167 | 128.5833 | 132.25 | 122.2167 | 115.9554 | 115.5 | 128.0667 | 127.3333 | 120.615 | 125.9333 | 127 | 129.0167 | 139.0213 | 130.3333 | 127.6059
| Median |  | 129.9825 | 126.925 | 121 | 125.25 | 123 | 125.25 | 129 | 132 | 122.45 | 115.7788 | 115.5 | 127.3 | 128.5 | 120.44 | 126.05 | 125.5 | 129.75 | 138.7832 | 130.5 | 129.4158
| Std Dev. |  | 0.56474 | 0.59365 | 1.6733 | 2.2749 | 1.8819 | 4.306 | 2.6724 | 1.6047 | 3.4539 | 3.2867 | 3.1464 | 1.7705 | 2.9439 | 2.6019 | 1.4868 | 5.3666 | 2.8188 | 2.601 | 2.3381 | 3.6466
| Rel.Std.Dev. |  | 0.4339% | 0.4677% | 1.383% | 1.816% | 1.542% | 3.393% | 2.078% | 1.213% | 2.826% | 2.834% | 2.724% | 1.382% | 2.312% | 2.157% | 1.181% | 4.226% | 2.185% | 1.871% | 1.794% | 2.858%
| PDM3 |  | 2.74% | 0.2085% | -4.474% | -1.119% | -3.619% | 0.1967% | 1.512% | 4.407% | -3.514% | -8.457% | -8.816% | 1.105% | 0.5256% | -4.778% | -0.5796% | 0.2625% | 1.855% | 9.753% | 2.894% | 0.7408%
| Z-Score (Absolute) |  | 0.82 | 0.01 | 1.54 | 0.44 | 1.26 | 0.01 | 0.42 | 1.36 | 1.23 | 2.84 | 2.96 | 0.28 | 0.09 | 1.64 | 0.27 | 0.01 | 0.53 | 3.11 | 0.87 | 0.17
| Analytical results for Re in OREAS 931b (Certified Value 0.002 ppm) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 24 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | < 0.05 | 0.003 | < 0.05 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.1 | < 0.002 | 0.002 | 0.004 | 0.002 | < 0.005 | 0.002 | < 0.002 | 0.0069 | [6 empty] | 1
| 1 | 2 | < 0.05 | 0.003 | < 0.05 | 0.002 | 0.002 | 0.002 | < 0.002 | 0.002 | < 0.1 | < 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.005 | 0.002 | < 0.002 | 0.0063 | [6 empty] | 0
| 1 | 3 | < 0.05 | 0.003 | < 0.05 | 0.003 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.1 | < 0.002 | < 0.002 | 0.004 | < 0.002 | < 0.005 | 0.003 | < 0.002 | 0.0069 | [6 empty] | 16
| 1 | 4 | < 0.05 | 0.003 | < 0.05 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | < 0.1 | < 0.002 | 0.002 | 0.005 | 0.003 | < 0.005 | 0.004 | < 0.002 | 0.0076 | [6 empty] | 0.0020185
| 1 | 5 | < 0.05 | < 0.002 | < 0.05 | 0.002 | 0.003 | < 0.002 | < 0.002 | 0.002 | < 0.1 | < 0.002 | 0.002 | 0.005 | < 0.002 | < 0.005 | 0.003 | < 0.002 | 0.0067 | [6 empty] | 44
| 1 | 6 | < 0.05 | 0.003 | < 0.05 | 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | < 0.1 | < 0.002 | < 0.002 | 0.005 | 0.003 | < 0.005 | 0.003 | < 0.002 | 0.0079
| Mean |  |  | 0.003 |  | 0.0021667 | 0.0021667 | 0.002 | 0.002 | 0.002 |  |  | 0.002 | 0.0043333 | 0.0026667 |  | 0.0028333 |  | 0.00705
| Median |  |  | 0.003 |  | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 |  |  | 0.002 | 0.0045 | 0.003 |  | 0.003 |  | 0.0069
| Std Dev. |  |  | 0 |  | 0.00040825 | 0.00040825 | 0 | 0 | 0 |  |  | 0 | 0.0008165 | 0.00057735 |  | 0.00075277 |  | 0.00059245
| Rel.Std.Dev. |  |  | 0% |  | 18.84% | 18.84% | 0% | 0% | 0% |  |  | 0% | 18.84% | 21.65% |  | 26.57% |  | 8.404%
| PDM3 |  |  | 48.62% |  | 7.339% | 7.339% | -0.9174% | -0.9174% | -0.9174% |  |  | -0.9174% | 114.7% | 32.11% |  | 40.37% |  | 249.3%
| Z-Score (Absolute) |  | 23.1 | 0.51 | 23.1 | 0 | 0 | 0.51 | 0.67 | 0.34 | 48.38 | 1.18 | 0.67 | 2.19 | 0.34 | 0.34 | 0.67 | 1.18 | 4.94
| Analytical results for S in OREAS 931b (Certified Value 4.3 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 4.34 | 4.4406 | 4.55 | 4.13 | 4.26 | 4.22 | 4.46 | 4.32 | 4.22 | 4.26 | 4.489 | 4.49 | 4 | 4.08 | 3.7 | 4.54 | 4.4 | 4.0365 | 4.096 | 4.24 | 4.15 | 3.6606 |  | 1
| 1 | 2 | 4.35 | 4.3964 | 4.48 | 4.15 | 4.34 | 4.16 | 4.44 | 4.22 | 4.25 | 4.44 | 4.613 | 4.54 | 4.1 | 4.19 | 3.7 | 4.5 | 4.14 | 3.8502 | 4.148 | 4.04 | 4.38 | 3.6085 |  | 0
| 1 | 3 | 4.37 | 4.4856 | 4.61 | 4.24 | 4.29 | 4.15 | 4.39 | 4.19 | 4.31 | 4.17 | 4.814 | 4.46 | 4 | 4.15 | 3.7 | 4.5 | 4.46 | 4.1589 | 4.02 | 4.18 | 4.27 | 3.7689 |  | 16
| 1 | 4 | 4.34 | 4.3817 | 4.62 | 4.2 | 4.26 | 4.18 | 4.33 | 4.19 | 4.26 | 4.32 | 4.859 | 4.47 | 4.1 | 4.3 | 3.5 | 4.54 | 4.42 | 4.0077 | 4.078 | 4.22 | 4.36 | 3.5524 |  | 4.2985
| 1 | 5 | 4.34 | 4.4266 | 4.57 | 4.16 | 4.36 | 4.19 | 4.22 | 4.21 | 4.23 | 4.36 | > 5 | 4.51 | 4 | 4.21 | 3.6 | 4.55 | 4.43 | 3.9051 | 4.114 | 4.2 | 4.17 | 3.6918 |  | 45
| 1 | 6 | 4.33 | 4.5075 | 4.54 | 4.2 | 4.33 | 4.18 | 4.28 | 4.22 | 4.21 | 4.1 | > 5 | 4.47 | 4 | 4.16 | 3.8 | 4.58 | 4.48 | 4.167 | 4.102 | 4.27 | 4.35 | 3.6405
| Mean |  | 4.345 | 4.4397 | 4.5617 | 4.18 | 4.3067 | 4.18 | 4.3533 | 4.225 | 4.2467 | 4.275 | 4.6937 | 4.49 | 4.0333 | 4.1817 | 3.6667 | 4.535 | 4.3883 | 4.0209 | 4.093 | 4.1917 | 4.28 | 3.6538
| Median |  | 4.34 | 4.4336 | 4.56 | 4.18 | 4.31 | 4.18 | 4.36 | 4.215 | 4.24 | 4.29 | 4.7135 | 4.48 | 4 | 4.175 | 3.7 | 4.54 | 4.425 | 4.0221 | 4.099 | 4.21 | 4.31 | 3.6506
| Std Dev. |  | 0.013784 | 0.04924 | 0.051153 | 0.040497 | 0.04274 | 0.024495 | 0.093737 | 0.048477 | 0.036148 | 0.12518 | 0.17341 | 0.030332 | 0.05164 | 0.073052 | 0.10328 | 0.030822 | 0.12497 | 0.12912 | 0.042685 | 0.080602 | 0.1004 | 0.073821
| Rel.Std.Dev. |  | 0.3172% | 1.109% | 1.121% | 0.9688% | 0.9924% | 0.586% | 2.153% | 1.147% | 0.8512% | 2.928% | 3.694% | 0.6755% | 1.28% | 1.747% | 2.817% | 0.6796% | 2.848% | 3.211% | 1.043% | 1.923% | 2.346% | 2.02%
| PDM3 |  | 1.081% | 3.285% | 6.121% | -2.758% | 0.1892% | -2.758% | 1.275% | -1.711% | -1.207% | -0.5475% | 9.194% | 4.454% | -6.17% | -2.719% | -14.7% | 5.501% | 2.089% | -6.459% | -4.782% | -2.486% | -0.4312% | -15%
| Z-Score (Absolute) |  | 0.47 | 1.12 | 1.97 | 0.67 | 0.2 | 0.67 | 0.52 | 0.36 | 0.21 | 0.02 | 2.88 | 1.47 | NA | 0.66 | NA | 1.78 | 0.77 | 1.77 | 1.28 | 0.59 | 0.02 | 4.31
| Analytical results for Sb in OREAS 931b (Certified Value 4.22 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | < 5 | 4.2171 | 4.47 | 3.9 | 3.93 | 3.78 | 4.16 | 4.21 | 4.23 | 3.85 | < 5 | 4.58 | 2.91 | 4.29 | 4.27 | 3.3 | 2.7 | 3.9 | 4.2614 | 4.32 | < 5 | 4.2701 |  | 1
| 1 | 2 | < 5 | 4.3663 | 4.33 | 4 | 3.95 | 4.03 | 4.17 | 4.35 | 4.41 | 3.83 | < 5 | 4.57 | 3.05 | 4.28 | 4.36 | 3.3 | 1.8 | 3.8 | 4.3822 | 4.45 | 6 | 4.1355 |  | 16
| 1 | 3 | < 5 | 4.3283 | 4.43 | 3.9 | 3.9 | 4.18 | 3.93 | 4.15 | 4.55 | 3.31 | < 5 | 4.83 | 3.1 | 4.17 | 4.14 | 3.6 | 1.3 | 4.2 | 4.4191 | 4.55 | < 5 | 4.2676 |  | 16
| 1 | 4 | < 5 | 4.272 | 4.34 | 3.6 | 3.94 | 4.22 | 3.99 | 4.42 | 4.35 | 4.02 | < 5 | 4.97 | 2.93 | 4.26 | 4.39 | 3.3 | 1.3 | 4 | 4.645 | 4.24 | 6 | 4.3921 |  | 4.2156
| 1 | 5 | < 5 | 4.3578 | 4.33 | 3.6 | 4.04 | 4.18 | 3.92 | 4.22 | 4.61 | 3.83 | < 5 | 4.91 | 2.92 | 4.28 | 4.25 | 3.2 | 1.6 | 4.2 | 4.2082 | 4.02 | 9 | 4.1358 |  | 46
| 1 | 6 | < 5 | 4.1702 | 4.34 | 4 | 3.99 | 4.25 | 3.83 | 4.34 | 4.34 | 4.05 | < 5 | 5.02 | 2.98 | 4.14 | 4.17 | 3.3 | 1.3 | 4 | 4.2553 | 4.43 | < 5 | 4.1761
| Mean |  |  | 4.2853 | 4.3733 | 3.8333 | 3.9583 | 4.1067 | 4 | 4.2817 | 4.415 | 3.815 |  | 4.8133 | 2.9817 | 4.2367 | 4.2633 | 3.3333 | 1.6667 | 4.0167 | 4.3619 | 4.335 | 7 | 4.2295
| Median |  |  | 4.3001 | 4.34 | 3.9 | 3.945 | 4.18 | 3.96 | 4.28 | 4.38 | 3.84 |  | 4.87 | 2.955 | 4.27 | 4.26 | 3.3 | 1.45 | 4 | 4.3218 | 4.375 | 6 | 4.2218
| Std Dev. |  |  | 0.07968 | 0.060882 | 0.18619 | 0.049565 | 0.17705 | 0.1377 | 0.10342 | 0.14167 | 0.26606 |  | 0.19521 | 0.077825 | 0.064704 | 0.099532 | 0.13663 | 0.5465 | 0.16021 | 0.16056 | 0.18812 | 1.7321 | 0.10001
| Rel.Std.Dev. |  |  | 1.859% | 1.392% | 4.857% | 1.252% | 4.311% | 3.442% | 2.416% | 3.209% | 6.974% |  | 4.056% | 2.61% | 1.527% | 2.335% | 4.099% | 32.79% | 3.989% | 3.681% | 4.34% | 24.74% | 2.364%
| PDM3 |  |  | 1.652% | 3.741% | -9.068% | -6.103% | -2.585% | -5.115% | 1.567% | 4.729% | -9.503% |  | 14.18% | -29.27% | 0.4991% | 1.132% | -20.93% | -60.46% | -4.72% | 3.469% | 2.832% | 66.05% | 0.3299%
| Z-Score (Absolute) |  | 4.25 | 0.47 | 0.7 | 0.72 | 0.39 | 0 | 0.28 | 0.46 | 0.81 | 0.77 | 4.25 | 1.87 | 2.97 | 0.34 | 0.41 | 2.04 | 6.45 | 0.24 | 0.67 | 0.6 | NA | 0.32
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 931b (Certified Value 10.1 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 10 | 9.9984 | 10.5 | 9 | 10 | 9.6 | 11.3 | 10.7 | 10 | 7.1 | 9.8908 | 9.8 | 10 | 10.8 | 9.6 | 10 | 10 | 10.3 | 11.7005 | 8 | 10 | 10 | 9.5296 | 1
| 1 | 2 | 9 | 10.105 | 10.5 | 9 | 9.9 | 9.4 | 11.2 | 10.6 | 10.2 | 7.5 | 9.7381 | 10.1 | 10 | 10.8 | 10.2 | 10 | 10 | 10.2 | 11.2603 | 8.5 | 10 | 9.9 | 9.5592 | 0
| 1 | 3 | 10 | 10.2135 | 10.4 | 9 | 9.9 | 9.7 | 10.2 | 10.1 | 10.2 | 7.1 | 9.1369 | 11.2 | 10.1 | 10.1 | 9.7 | 10 | 10 | 10.3 | 12.5084 | 9 | 10 | 10 | 9.8472 | 16
| 1 | 4 | 10 | 10.1759 | 10.5 | 9 | 9.4 | 9.5 | 10.5 | 10.7 | 10.4 | 8.1 | 9.9709 | 11.6 | 10.1 | 10.9 | 10.4 | 10 | 11 | 10 | 12.3144 | 8.6 | 10 | 10.1 | 9.7673 | 10.0955
| 1 | 5 | 9 | 9.9785 | 10.5 | 9 | 9.6 | 9.6 | 10.1 | 10.1 | 9.9 | 7.6 | 9.5107 | 11.8 | 10.1 | 10.8 | 10.1 | 10 | 11 | 10.3 | 11.59 | 9.2 | 10 | 10.5 | 9.5394 | 47
| 1 | 6 | 9 | 10.1732 | 10.4 | 9 | 9.6 | 9.7 | 10.4 | 10.5 | 10.7 | 6.7 | 9.4707 | 11.3 | 10.1 | 10.6 | 10 | 11 | 11 | 10.4 | 12.0277 | 9 | 10 | 10.3 | 9.7179
| Mean |  | 9.5 | 10.1074 | 10.4667 | 9 | 9.7333 | 9.5833 | 10.6167 | 10.45 | 10.2333 | 7.35 | 9.6197 | 10.9667 | 10.0667 | 10.6667 | 10 | 10.1667 | 10.5 | 10.25 | 11.9002 | 8.7167 | 10 | 10.1333 | 9.6601
| Median |  | 9.5 | 10.1391 | 10.5 | 9 | 9.75 | 9.6 | 10.45 | 10.55 | 10.2 | 7.3 | 9.6244 | 11.25 | 10.1 | 10.8 | 10.05 | 10 | 10.5 | 10.3 | 11.8641 | 8.8 | 10 | 10.05 | 9.6385
| Std Dev. |  | 0.54772 | 0.098781 | 0.05164 | 0 | 0.23381 | 0.1169 | 0.51153 | 0.28107 | 0.28752 | 0.48888 | 0.3092 | 0.82138 | 0.05164 | 0.29439 | 0.30332 | 0.40825 | 0.54772 | 0.13784 | 0.46973 | 0.44008 | 0 | 0.22509 | 0.13536
| Rel.Std.Dev. |  | 5.766% | 0.9773% | 0.4934% | 0% | 2.402% | 1.22% | 4.818% | 2.69% | 2.81% | 6.651% | 3.214% | 7.49% | 0.513% | 2.76% | 3.033% | 4.016% | 5.216% | 1.345% | 3.947% | 5.049% | 0% | 2.221% | 1.401%
| PDM3 |  | -5.898% | 0.1187% | 3.677% | -10.85% | -3.587% | -5.073% | 5.163% | 3.512% | 1.366% | -27.19% | -4.713% | 8.63% | -0.2851% | 5.658% | -0.9455% | 0.7054% | 4.007% | 1.531% | 17.88% | -13.66% | -0.9455% | 0.3752% | -4.312%
| Z-Score (Absolute) |  | NA | 0.02 | 0.55 | NA | 0.62 | 0.85 | 0.79 | 0.52 | 0.18 | 4.41 | 0.8 | 1.35 | 0.09 | 0.87 | 0.19 | NA | NA | 0.21 | 2.83 | 2.23 | NA | 0.02 | 0.73
Results from laboratories 1, 4, 16, 17 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 931b (Certified Value 45.9 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 24 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 47 | 50.0063 | 45.9 | 49 | 44 | 54 | 45 | 52 | 46 | 37 | 37 | 45.6 | 50.6 | 45 | 46 | 44.2 | 48 | 37 | 40.4529 |  |  |  |  | 1
| 1 | 2 | 46 | 50.1216 | 47.1 | 47 | 46 | 56 | 45 | 46 | 46 | 40 | 36 | 45.1 | 48.6 | 46.9 | 45 | 44.1 | 49 | 39 | 41.7769 |  |  |  |  | 0
| 1 | 3 | 47 | 50.7744 | 45.3 | 49 | 47 | 47 | 44 | 45 | 49 | 38 | 36 | 45.6 | 47.2 | 43.7 | 46 | 45.6 | 48 | 40 | 39.5699 |  |  |  |  | 16
| 1 | 4 | 47 | 49.7131 | 48.9 | 49 | 44 | 47 | 42 | 45 | 48 | 39 | 37 | 44.1 | 50.4 | 45.9 | 49 | 44.1 | 46 | 39 | 39.3018 |  |  |  |  | 45.9473
| 1 | 5 | 46 | 51.1262 | 46.9 | 48 | 45 | 47 | 42 | 47 | 46 | 36 | 35 | 44.8 | 50.9 | 44.9 | 46 | 46.5 | 48 | 40 | 36.9915 |  |  |  |  | 48
| 1 | 6 | 46 | 49.2725 | 45.6 | 48 | 46 | 48 | 42 | 45 | 50 | 38 | 36 | 43.9 | 50.4 | 45 | 50 | 46.3 | 45 | 38 | 37.8355
| Mean |  | 46.5 | 50.169 | 46.6167 | 48.3333 | 45.3333 | 49.8333 | 43.3333 | 46.6667 | 47.5 | 38 | 36.1667 | 44.85 | 49.6833 | 45.2333 | 47 | 45.1333 | 47.3333 | 38.8333 | 39.3214
| Median |  | 46.5 | 50.064 | 46.4 | 48.5 | 45.5 | 47.5 | 43 | 45.5 | 47 | 38 | 36 | 44.95 | 50.4 | 45 | 46 | 44.9 | 48 | 39 | 39.4359
| Std Dev. |  | 0.54772 | 0.68161 | 1.3273 | 0.8165 | 1.2111 | 4.0702 | 1.5055 | 2.7325 | 1.7607 | 1.4142 | 0.75277 | 0.7287 | 1.4621 | 1.0764 | 2 | 1.1361 | 1.5055 | 1.169 | 1.7325
| Rel.Std.Dev. |  | 1.178% | 1.359% | 2.847% | 1.689% | 2.671% | 8.168% | 3.474% | 5.855% | 3.707% | 3.722% | 2.081% | 1.625% | 2.943% | 2.38% | 4.255% | 2.517% | 3.181% | 3.01% | 4.406%
| PDM3 |  | 1.203% | 9.188% | 1.457% | 5.193% | -1.336% | 8.458% | -5.689% | 1.566% | 3.379% | -17.3% | -21.29% | -2.388% | 8.131% | -1.554% | 2.291% | -1.771% | 3.017% | -15.48% | -14.42%
| Z-Score (Absolute) |  | 0 | 1.5 | 0.05 | 0.75 | 0.48 | 1.36 | 1.29 | 0.07 | 0.41 | 3.47 | 4.22 | 0.67 | 1.3 | 0.52 | 0.2 | 0.56 | 0.34 | 3.13 | 2.93
| Analytical results for Sm in OREAS 931b (Certified Value 5.77 ppm) | [24 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  |  | 1
| No | No | 01: 03; 02: 11; 03: 12; 04: 13; 05: 15; 06: 17; 07: 22; 08: 24 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*OES; 08: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 01: 5.85; 02: 5.2415; 03: 6; 04: 5.1; 05: 5.65; 06: 5.3; 07: 5; 08: 6.48 |  |  |  |  | 1
| 1 | 2 | 01: 5.94; 02: 5.2162; 03: 6.06; 04: 5.2; 05: 5.79; 06: 5.3; 07: 5; 08: 6.2387 |  |  |  |  | 17
| 1 | 3 | 01: 6.01; 02: 5.0686; 03: 6.36; 04: 5.3; 05: 5.71; 06: 5.4; 07: 5; 08: 6.715 |  |  |  |  | 16
| 1 | 4 | 01: 5.9; 02: 5.458; 03: 6.47; 04: 5; 05: 5.73; 06: 6; 07: 5; 08: 6.374 |  |  |  |  | 5.7732
| 1 | 5 | 01: 5.88; 02: 5.2755; 03: 6.37; 04: 5.1; 05: 5.97; 06: 6.3; 07: 5; 08: 6.3343 |  |  |  |  | 49
| 1 | 6 | 01: 5.83; 02: 5.2029; 03: 6.7; 04: 5; 05: 5.77; 06: 5.9; 07: 5; 08: 6.3394
| Mean |  | 01: 5.9017; 02: 5.2438; 03: 6.3267; 04: 5.1167; 05: 5.77; 06: 5.7; 07: 5; 08: 6.4136
| Median |  | 01: 5.89; 02: 5.2289; 03: 6.365; 04: 5.1; 05: 5.75; 06: 5.65; 07: 5; 08: 6.3567
| Std Dev. |  | 01: 0.065549; 02: 0.12651; 03: 0.26105; 04: 0.1169; 05: 0.10954; 06: 0.42426; 07: 0; 08: 0.16686
| Rel.Std.Dev. |  | 01: 1.111%; 02: 2.413%; 03: 4.126%; 04: 2.285%; 05: 1.899%; 06: 7.443%; 07: 0%; 08: 2.602%
| PDM3 |  | 01: 2.226%; 02: -9.169%; 03: 9.588%; 04: -11.37%; 05: -0.05457%; 06: -1.267%; 07: -13.39%; 08: 11.09%
| Z-Score (Absolute) |  | 01: 0.17; 02: 0.67; 03: 0.71; 04: 0.84; 05: 0; 06: 0.09; 07: NA; 08: 0.82
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 931b (Certified Value 37.4 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 41 | 37.8517 | 37.1 | 33.6 | 38 | 33.7 | 37.3 | 37 | 37.1 | 36.3 | 29.919 | 35.3 | 37.4 | 38.1 | 37.1 | 39.1 | 37 | 37.5 | 60.939 | 37.6 | 15 | 36 | 28.2351 | 1
| 1 | 2 | 38 | 37.9356 | 37.2 | 33.1 | 37.7 | 35.3 | 38.5 | 36.9 | 38 | 38.2 | 31.9732 | 35.7 | 37.3 | 38.4 | 38.1 | 38.3 | 36 | 38.6 | 60.3691 | 38.8 | 15 | 36 | 28.033 | 8
| 1 | 3 | 40 | 37.0334 | 37.1 | 33.9 | 38.1 | 36 | 35.2 | 35.6 | 38.2 | 36.1 | 31.3045 | 37 | 37.8 | 37.5 | 36 | 40.9 | 37 | 38.4 | 63.8696 | 38.6 | 14 | 36 | 28.3326 | 16
| 1 | 4 | 39 | 37.8062 | 37.4 | 33.4 | 37.9 | 35 | 35.6 | 38.1 | 37.9 | 37.6 | 32.7962 | 37.2 | 36.6 | 39.1 | 38.5 | 37.8 | 38 | 36.6 | 64.3637 | 37.7 | 17 | 37 | 28.2308 | 37.4068
| 1 | 5 | 40 | 38.3118 | 36.8 | 34.1 | 39.4 | 36 | 35.8 | 36.5 | 38.2 | 38.5 |  | 37.3 | 36.5 | 39 | 37.5 | 39.4 | 40 | 38 | 63.8677 | 38.8 | 15 | 40 | 28.5358 | 50
| 1 | 6 | 40 | 36.9353 | 36.9 | 33.8 | 38.2 | 36 | 34.9 | 37.6 | 37.3 | 36.5 | 32.3083 | 38.2 | 36.6 | 38 | 36 | 40.2 | 39 | 37.5 | 64.3971 | 37.5 | 14 | 38 | 28.4335
| Mean |  | 39.6667 | 37.6457 | 37.0833 | 33.65 | 38.2167 | 35.3333 | 36.2167 | 36.95 | 37.7833 | 37.2 | 31.6602 | 36.7833 | 37.0333 | 38.35 | 37.2 | 39.2833 | 37.8333 | 37.7667 | 62.9677 | 38.1667 | 15 | 37.1667 | 28.3001
| Median |  | 40 | 37.8289 | 37.1 | 33.7 | 38.05 | 35.65 | 35.7 | 36.95 | 37.95 | 37.05 | 31.9732 | 37.1 | 36.95 | 38.25 | 37.3 | 39.25 | 37.5 | 37.75 | 63.8686 | 38.15 | 15 | 36.5 | 28.2838
| Std Dev. |  | 1.0328 | 0.54325 | 0.2137 | 0.36194 | 0.6047 | 0.90701 | 1.3934 | 0.86891 | 0.47081 | 1.0354 | 1.1143 | 1.0834 | 0.53914 | 0.61563 | 1.0469 | 1.1548 | 1.472 | 0.72847 | 1.8157 | 0.62823 | 1.0954 | 1.6021 | 0.17595
| Rel.Std.Dev. |  | 2.604% | 1.443% | 0.5763% | 1.076% | 1.582% | 2.567% | 3.848% | 2.352% | 1.246% | 2.783% | 3.52% | 2.945% | 1.456% | 1.605% | 2.814% | 2.94% | 3.891% | 1.929% | 2.884% | 1.646% | 7.303% | 4.311% | 0.6217%
| PDM3 |  | 6.041% | 0.6385% | -0.8648% | -10.04% | 2.165% | -5.543% | -3.182% | -1.221% | 1.007% | -0.5529% | -15.36% | -1.667% | -0.9984% | 2.521% | -0.5529% | 5.016% | 1.14% | 0.962% | 68.33% | 2.031% | -59.9% | -0.642% | -24.34%
| Z-Score (Absolute) |  | NA | 0.31 | 0.08 | 2.48 | 0.71 | 1.3 | 0.69 | 0.17 | 0.41 | 0 | 3.86 | 0.29 | 0.12 | 0.8 | 0 | 1.45 | NA | 0.4 | 17.97 | 0.67 | NA | NA | 6.21
Results from laboratories 1, 17, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 931b (Certified Value 46.6 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 60 | 50.1849 | 49.81 | 43.7 | 47.4 | 45.8 | 48.8 | 48.1 | 45.9 | 34.6 | 47.2327 | 54.5 | 45 | 45.7 | 46.39 | 42 | 40.4 | 47.8 | 45.6532 | 44.7 | 51 | 48.1 | 59.8019 | 1
| 1 | 2 | 60 | 49.7256 | 49.8 | 44.4 | 48.6 | 46.1 | 48 | 49.3 | 46.4 | 44 | 46.008 | 55.8 | 46 | 45.1 | 48.95 | 45 | 41.1 | 50.6 | 50.5335 | 43.7 | 51 | 47.1 | 59.3191 | 0
| 1 | 3 | 59 | 49.8008 | 50.1 | 44.9 | 47.9 | 43.6 | 45.6 | 46.6 | 47.5 | 34.7 | 44.2552 | 58 | 45 | 42.4 | 46.11 | 45 | 40.8 | 49.5 | 50.0045 | 43.7 | 50 | 46.3 | 61.6044 | 16
| 1 | 4 | 60 | 49.9421 | 49.2 | 44.9 | 46.6 | 42.9 | 46.5 | 48.5 | 46.6 | 43.6 | 47.7286 | 59.6 | 46 | 45.3 | 49.51 | 44 | 39.2 | 46.7 | 48.1801 | 43.2 | 50 | 47.3 | 61.2609 | 46.6102
| 1 | 5 | 61 | 49.8952 | 49.98 | 46.4 | 48.8 | 44.1 | 45.6 | 47.8 | 47.4 | 42.9 | 45.8428 | 65.1 | 45 | 45.8 | 47 | 45 | 44.1 | 48.9 | 45.1998 | 44.1 | 51 | 45.7 | 59.2055 | 51
| 1 | 6 | 60 | 50.1671 | 49.9 | 46.5 | 47.4 | 43.2 | 48 | 48.7 | 46.6 | 34.7 | 45.5967 | 62.2 | 46 | 44.6 | 47.18 | 44 | 42.7 | 47.7 | 42.4979 | 44.2 | 52 | 47.7 | 60.5947
| Mean |  | 60 | 49.9526 | 49.7983 | 45.1333 | 47.7833 | 44.2833 | 47.0833 | 48.1667 | 46.7333 | 39.0833 | 46.1107 | 59.2 | 45.5 | 44.8167 | 47.5233 | 44.1667 | 41.3833 | 48.5333 | 47.0115 | 43.9333 | 50.8333 | 47.0333 | 60.2977
| Median |  | 60 | 49.9187 | 49.855 | 44.9 | 47.65 | 43.85 | 47.25 | 48.3 | 46.6 | 38.8 | 45.9254 | 58.8 | 45.5 | 45.2 | 47.09 | 44.5 | 40.95 | 48.35 | 46.9167 | 43.9 | 51 | 47.2 | 60.1983
| Std Dev. |  | 0.63246 | 0.18868 | 0.31384 | 1.1112 | 0.82563 | 1.3556 | 1.3688 | 0.92448 | 0.6121 | 4.8512 | 1.2382 | 3.9764 | 0.54772 | 1.2608 | 1.3897 | 1.169 | 1.7475 | 1.4095 | 3.1058 | 0.5164 | 0.75277 | 0.89144 | 1.0119
| Rel.Std.Dev. |  | 1.054% | 0.3777% | 0.6302% | 2.462% | 1.728% | 3.061% | 2.907% | 1.919% | 1.31% | 12.41% | 2.685% | 6.717% | 1.204% | 2.813% | 2.924% | 2.647% | 4.223% | 2.904% | 6.606% | 1.175% | 1.481% | 1.895% | 1.678%
| PDM3 |  | 28.73% | 7.171% | 6.84% | -3.169% | 2.517% | -4.992% | 1.015% | 3.339% | 0.2642% | -16.15% | -1.072% | 27.01% | -2.382% | -3.848% | 1.959% | -5.242% | -11.21% | 4.126% | 0.861% | -5.743% | 9.061% | 0.9078% | 29.37%
| Z-Score (Absolute) |  | NA | 1.04 | 0.98 | 0.67 | 0.27 | 0.98 | 0.02 | 0.4 | 0.11 | 2.82 | 0.33 | 4.32 | NA | 0.79 | 0.17 | NA | 2.01 | 0.53 | 0.01 | 1.1 | NA | 0 | 4.71
Results from laboratories 1, 13, 16 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 931b (Certified Value 0.2 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  | 2
2
| 1 | 1 | < 5 | 0.23993 | 0.2 | 0.14 | 0.18 | 0.29 | 0.19 | 0.17 | 0.17 | 0.29 | 0.307 | 0.51 | 0.2 | 0.16 | 0.17 | 0.1 | 0.6 | 0.21 | 0.21906 | 0.12 | < 1 | 0.1665 |  | 1
| 1 | 2 | < 5 | 0.24679 | 0.21 | 0.13 | 0.16 | 0.32 | 0.19 | 0.17 | 0.2 | 0.27 | 0.3455 | 0.48 | 0.2 | 0.16 | 0.2 | 0.1 | 0.6 | 0.19 | 0.26367 | 0.12 | < 1 | 0.1687 |  | 10
| 1 | 3 | < 5 | 0.23531 | 0.23 | 0.13 | 0.12 | 0.25 | 0.2 | 0.2 | 0.17 | 0.26 | 0.3324 | 0.63 | 0.2 | 0.18 | 0.16 | 0.1 | 0.6 | 0.22 | 0.2508 | 0.12 | < 1 | 0.1345 |  | 16
| 1 | 4 | < 5 | 0.22254 | 0.25 | 0.13 | 0.18 | 0.26 | 0.19 | 0.17 | 0.18 | 0.29 | 0.4579 | 0.71 | 0.2 | 0.17 | 0.17 | 0.1 | 0.6 | 0.2 | 0.23598 | 0.12 | < 1 | 0.1582 |  | 0.19686
| 1 | 5 | < 5 | 0.2193 | 0.28 | 0.1 | 0.17 | 0.24 | 0.17 | 0.17 | 0.2 | 0.27 |  | 0.83 | 0.2 | 0.17 | 0.2 | 0.1 | 0.6 | 0.21 | 0.22872 | 0.11 | < 1 | 0.1593 |  | 52
| 1 | 6 | < 5 | 0.23965 | 0.31 | 0.13 | 0.17 | 0.25 | 0.19 | 0.18 | 0.24 | 0.34 | 0.3253 | 0.61 | 0.2 | 0.16 | 0.23 | 0.1 | 0.6 | 0.18 | 0.17575 | 0.12 | < 1 | 0.1466
| Mean |  |  | 0.23392 | 0.24667 | 0.12667 | 0.16333 | 0.26833 | 0.18833 | 0.17667 | 0.19333 | 0.28667 | 0.35362 | 0.62833 | 0.2 | 0.16667 | 0.18833 | 0.1 | 0.6 | 0.20167 | 0.229 | 0.11833 |  | 0.15563
| Median |  |  | 0.23748 | 0.24 | 0.13 | 0.17 | 0.255 | 0.19 | 0.17 | 0.19 | 0.28 | 0.3324 | 0.62 | 0.2 | 0.165 | 0.185 | 0.1 | 0.6 | 0.205 | 0.23235 | 0.12 |  | 0.15875
| Std Dev. |  |  | 0.010768 | 0.042269 | 0.013663 | 0.022509 | 0.030605 | 0.0098319 | 0.012111 | 0.026583 | 0.028752 | 0.059929 | 0.12937 | 3.0405e-17 | 0.008165 | 0.026394 | 1.5202e-17 | 0 | 0.01472 | 0.03053 | 0.0040825 |  | 0.012939
| Rel.Std.Dev. |  |  | 4.603% | 17.14% | 10.79% | 13.78% | 11.41% | 5.22% | 6.855% | 13.75% | 10.03% | 16.95% | 20.59% | 1.52e-14% | 4.899% | 14.01% | 1.52e-14% | 0% | 7.299% | 13.33% | 3.45% |  | 8.314%
| PDM3 |  |  | 18.83% | 25.3% | -35.66% | -17.03% | 36.31% | -4.331% | -10.26% | -1.791% | 45.62% | 79.63% | 219.2% | 1.596% | -15.34% | -4.331% | -49.2% | 204.8% | 2.442% | 16.32% | -39.89% |  | -20.94%
| Z-Score (Absolute) |  | 34.44 | 0.48 | 0.67 | 1.12 | 0.57 | 1 | 0.2 | 0.37 | 0.12 | 1.27 | 2.28 | 6.39 | NA | 0.52 | 0.2 | NA | NA | 0 | 0.41 | 1.25 | 4.47 | 0.69
Results from laboratories 13, 16 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 931b (Certified Value 0.59 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  | 1
| No | No | 01: 03; 02: 10; 03: 11; 04: 12; 05: 13; 06: 15; 07: 17; 08: 20; 09: 22 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*OES |  | 2
3
| 1 | 1 | 01: 0.56; 02: 0.58; 03: 0.608; 04: 0.6; 05: 0.6; 06: 0.52; 07: 0.7; 08: 0.56; 09: < 1 |  | 1
| 1 | 2 | 01: 0.55; 02: 0.62; 03: 0.5969; 04: 0.61; 05: 0.5; 06: 0.55; 07: 0.7; 08: 0.61; 09: < 1 |  | 0
| 1 | 3 | 01: 0.56; 02: 0.58; 03: 0.5822; 04: 0.66; 05: 0.6; 06: 0.5; 07: 0.7; 08: 0.58; 09: < 1 |  | 16
| 1 | 4 | 01: 0.55; 02: 0.64; 03: 0.623; 04: 0.66; 05: 0.5; 06: 0.54; 07: 0.7; 08: 0.58; 09: < 1 |  | 0.58842
| 1 | 5 | 01: 0.54; 02: 0.64; 03: 0.6088; 04: 0.66; 05: 0.5; 06: 0.54; 07: 0.7; 08: 0.6; 09: < 1 |  | 53
| 1 | 6 | 01: 0.55; 02: 0.58; 03: 0.6062; 04: 0.68; 05: 0.5; 06: 0.54; 07: 0.7; 08: 0.58; 09: < 1
| Mean |  | 01: 0.55167; 02: 0.60667; 03: 0.60418; 04: 0.645; 05: 0.53333; 06: 0.53167; 07: 0.7; 08: 0.585
| Median |  | 01: 0.55; 02: 0.6; 03: 0.6071; 04: 0.66; 05: 0.5; 06: 0.54; 07: 0.7; 08: 0.58
| Std Dev. |  | 01: 0.0075277; 02: 0.030111; 03: 0.013641; 04: 0.032094; 05: 0.05164; 06: 0.018348; 07: 1.2162e-16; 08: 0.017607
| Rel.Std.Dev. |  | 01: 1.365%; 02: 4.963%; 03: 2.258%; 04: 4.976%; 05: 9.682%; 06: 3.451%; 07: 1.737e-14%; 08: 3.01%
| PDM3 |  | 01: -6.246%; 02: 3.101%; 03: 2.679%; 04: 9.616%; 05: -9.362%; 06: -9.645%; 07: 18.96%; 08: -0.5811%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.44; 03: 0.39; 04: 1.21; 05: NA; 06: 1.08; 07: NA; 08: 0; 09: 1.72
Results from laboratories 13 and 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 931b (Certified Value 0.17 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  |  | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 13 | 0.16915 | < 0.2 | 0.2 | 0.17 | 0.15 | 0.19 | 0.18 | 0.15 | 0.19 | 0.1867 | 0.27 | 0.16 | 0.15 | < 0.2 | < 0.5 | 0.1 | 0.2 | < 0.1 | 0.17 | 0.1046 |  |  | 1
| 1 | 2 | 13 | 0.21021 | < 0.2 | 0.2 | 0.18 | 0.13 | 0.19 | 0.15 | 0.18 | 0.17 | 0.1489 | 0.2 | 0.16 | 0.19 | < 0.2 | < 0.5 | 0.1 | 0.2 | < 0.1 | 0.16 | 0.2267 |  |  | 21
| 1 | 3 | 13 | 0.15418 | < 0.2 | 0.2 | 0.17 | 0.17 | 0.19 | 0.15 | 0.18 | 0.16 | 0.1459 | 0.24 | 0.17 | 0.18 | < 0.2 | < 0.5 | 0.2 | 0.2 | < 0.1 | 0.19 | 0.1743 |  |  | 16
| 1 | 4 | 13 | 0.21178 | < 0.2 | 0.2 | 0.16 | 0.18 | 0.16 | 0.19 | 0.15 | 0.21 | 0.1283 | 0.26 | 0.16 | 0.2 | < 0.2 | < 0.5 | 0.1 | 0.2 | < 0.1 | 0.18 | 0.1395 |  |  | 0.171
| 1 | 5 | 13 | 0.16388 | < 0.2 | 0.2 | 0.19 | 0.19 | 0.17 | 0.17 | 0.16 | 0.15 | 0.1181 | 0.24 | 0.17 | 0.18 | < 0.2 | < 0.5 | 0.1 | 0.21 | < 0.1 | 0.16 | 0.1831 |  |  | 54
| 1 | 6 | 13 | 0.18409 | < 0.2 | 0.2 | 0.18 | 0.18 | 0.17 | 0.14 | 0.19 | 0.18 | 0.1756 | 0.25 | 0.16 | 0.13 | < 0.2 | < 0.5 | 0.1 | 0.18 | < 0.1 | 0.17 | 0.1133
| Mean |  | 13 | 0.18221 |  | 0.2 | 0.175 | 0.16667 | 0.17833 | 0.16333 | 0.16833 | 0.17667 | 0.15058 | 0.24333 | 0.16333 | 0.17167 |  |  | 0.11667 | 0.19833 |  | 0.17167 | 0.15692
| Median |  | 13 | 0.17662 |  | 0.2 | 0.175 | 0.175 | 0.18 | 0.16 | 0.17 | 0.175 | 0.1474 | 0.245 | 0.16 | 0.18 |  |  | 0.1 | 0.2 |  | 0.17 | 0.1569
| Std Dev. |  | 0 | 0.024307 |  | 3.0405e-17 | 0.010488 | 0.022509 | 0.013292 | 0.019664 | 0.017224 | 0.021602 | 0.026483 | 0.024221 | 0.005164 | 0.026394 |  |  | 0.040825 | 0.0098319 |  | 0.01169 | 0.046476
| Rel.Std.Dev. |  | 0% | 13.34% |  | 1.52e-14% | 5.993% | 13.51% | 7.453% | 12.04% | 10.23% | 12.23% | 17.59% | 9.954% | 3.162% | 15.38% |  |  | 34.99% | 4.957% |  | 6.81% | 29.62%
| PDM3 |  | 7502% | 6.555% |  | 16.96% | 2.337% | -2.536% | 4.286% | -4.485% | -1.562% | 3.312% | -11.94% | 42.3% | -4.485% | 0.3878% |  |  | -31.78% | 15.98% |  | 0.3878% | -8.238%
| Z-Score (Absolute) |  | NA | 0.8 | 4.59 | NA | 0.33 | 0.22 | 0.55 | 0.44 | 0.11 | 0.44 | 1.27 | 4.81 | 0.44 | 0.11 | 4.59 | 5.25 | NA | 1.86 | 7.88 | 0.11 | 0.86
Results from laboratories 4 and 17 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 931b (Certified Value 11.3 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 12.0002 | 11.14 | 9.69 | 11.8 | 11.4 | 11.9 | 11.5 | 10.65 | 10.3 | 10.595 | 11.65 | 10.1 | 12 | 10.85 | 11.2 | 6.5 | 11 | 14.4179 | 11.62 | 12 | 12.6364 |  |  | 1
| 1 | 2 | 12.0227 | 11.22 | 9.6 | 12.2 | 11.6 | 12.1 | 11.6 | 10.05 | 10.7 | 10.61 | 11.61 | 9 | 12.3 | 11.65 | 11.1 | 5.6 | 10.6 | 13.8744 | 11.82 | 12 | 12.8168 |  |  | 11
| 1 | 3 | 11.904 | 11.21 | 9.58 | 12 | 11 | 11 | 10.75 | 10.4 | 9.7 | 10.3394 | 12.49 | 9.6 | 12.2 | 11.13 | 11 | 7.3 | 11.3 | 14.3079 | 11.67 | 12 | 12.9887 |  |  | 16
| 1 | 4 | 11.9917 | 11.32 | 9.47 | 12.35 | 10.85 | 11.65 | 11.8 | 10.45 | 11.2 | 11.1915 | 12.79 | 12.2 | 12.6 | 11.77 | 11.3 | 7.4 | 10.6 | 13.7316 | 11.95 | 12 | 13.3389 |  |  | 11.2869
| 1 | 5 | 11.9681 | 11.49 | 9.31 | 12.3 | 10.7 | 11.55 | 11 | 10.5 | 10.9 | 10.9389 | 12.57 | 8.8 | 12.8 | 11.33 | 11.5 | 6.1 | 11.2 | 13.8705 | 11.64 | 12 | 12.6155 |  |  | 55
| 1 | 6 | 11.9176 | 11.25 | 9.53 | 12.65 | 10.75 | 11.25 | 11.6 | 11.05 | 10.4 | 10.7677 | 12.85 | 9.7 | 12.6 | 11.37 | 11.6 | 6.1 | 10.9 | 13.9394 | 11.78 | 12 | 12.9523
| Mean |  | 11.9674 | 11.2717 | 9.53 | 12.2167 | 11.05 | 11.575 | 11.375 | 10.5167 | 10.5333 | 10.7404 | 12.3267 | 9.9 | 12.4167 | 11.35 | 11.2833 | 6.5 | 10.9333 | 14.0236 | 11.7467 | 12 | 12.8914
| Median |  | 11.9799 | 11.235 | 9.555 | 12.25 | 10.925 | 11.6 | 11.55 | 10.475 | 10.55 | 10.6889 | 12.53 | 9.65 | 12.45 | 11.35 | 11.25 | 6.3 | 10.95 | 13.9069 | 11.725 | 12 | 12.8846
| Std Dev. |  | 0.047389 | 0.12189 | 0.13038 | 0.29439 | 0.36878 | 0.40589 | 0.40712 | 0.32813 | 0.52409 | 0.29752 | 0.55605 | 1.2231 | 0.29944 | 0.33633 | 0.23166 | 0.71833 | 0.29439 | 0.27362 | 0.12738 | 0 | 0.26838
| Rel.Std.Dev. |  | 0.396% | 1.081% | 1.368% | 2.41% | 3.337% | 3.507% | 3.579% | 3.12% | 4.976% | 2.77% | 4.511% | 12.35% | 2.412% | 2.963% | 2.053% | 11.05% | 2.693% | 1.951% | 1.084% | 0% | 2.082%
| PDM3 |  | 6.029% | -0.135% | -15.57% | 8.238% | -2.099% | 2.553% | 0.7805% | -6.824% | -6.676% | -4.842% | 9.212% | -12.29% | 10.01% | 0.559% | -0.03161% | -42.41% | -3.133% | 24.25% | 4.073% | 6.318% | 14.22%
| Z-Score (Absolute) |  | 0.61 | 0.04 | 1.68 | 0.85 | 0.25 | 0.24 | 0.05 | 0.75 | 0.74 | 0.54 | 0.95 | 1.33 | 1.03 | 0.03 | 0.03 | 4.53 | 0.36 | 2.55 | 0.4 | NA | 1.48
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 931b (Certified Value 0.15 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 |  | wt.%
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.13 | 0.17405 | 0.1604 | 0.139 | 0.155 | 0.169 | 0.161 | 0.154 | 0.156 | 0.17 | 0.137 | 0.123 | 0.14 | 0.1468 | 0.132 | 0.272 | 0.17 | 0.2254 | 0.143 | 0.252 | 0.15 | 0.14614 |  | 1
| 1 | 2 | 0.13 | 0.16847 | 0.164 | 0.135 | 0.154 | 0.166 | 0.163 | 0.152 | 0.161 | 0.17 | 0.137 | 0.123 | 0.14 | 0.1465 | 0.134 | 0.275 | 0.16 | 0.2346 | 0.138 | 0.252 | 0.14 | 0.14735 |  | 0
| 1 | 3 | 0.13 | 0.1731 | 0.167 | 0.139 | 0.137 | 0.163 | 0.167 | 0.158 | 0.16 | 0.16 | 0.139 | 0.121 | 0.14 | 0.1474 | 0.133 | 0.278 | 0.17 | 0.2272 | 0.135 | 0.246 | 0.14 | 0.14934 |  | 16
| 1 | 4 | 0.13 | 0.16322 | 0.1748 | 0.141 | 0.159 | 0.161 | 0.159 | 0.152 | 0.159 | 0.17 | 0.139 | 0.119 | 0.14 | 0.1482 | 0.134 | 0.281 | 0.17 | 0.2248 | 0.136 | 0.246 | 0.15 | 0.14866 |  | 0.15031
| 1 | 5 | 0.13 | 0.16237 | 0.1742 | 0.141 | 0.155 | 0.163 | 0.151 | 0.155 | 0.161 | 0.16 | 0.148 | 0.122 | 0.14 | 0.1515 | 0.136 | 0.263 | 0.17 | 0.2272 | 0.139 | 0.252 | 0.14 | 0.14534 |  | 56
| 1 | 6 | 0.13 | 0.1728 | 0.1869 | 0.143 | 0.156 | 0.161 | 0.157 | 0.154 | 0.161 | 0.17 | 0.15 | 0.123 | 0.14 | 0.1513 | 0.138 | 0.267 | 0.17 | 0.2286 | 0.141 | 0.252 | 0.15 | 0.14904
| Mean |  | 0.13 | 0.169 | 0.17122 | 0.13967 | 0.15267 | 0.16383 | 0.15967 | 0.15417 | 0.15967 | 0.16667 | 0.14167 | 0.12183 | 0.14 | 0.14862 | 0.1345 | 0.27267 | 0.16833 | 0.22797 | 0.13867 | 0.25 | 0.145 | 0.14765
| Median |  | 0.13 | 0.17064 | 0.1706 | 0.14 | 0.155 | 0.163 | 0.16 | 0.154 | 0.1605 | 0.17 | 0.139 | 0.1225 | 0.14 | 0.1478 | 0.134 | 0.2735 | 0.17 | 0.2272 | 0.1385 | 0.252 | 0.145 | 0.14801
| Std Dev. |  | 0 | 0.0051855 | 0.0095349 | 0.0027325 | 0.0078655 | 0.0031252 | 0.005465 | 0.0022286 | 0.0019664 | 0.005164 | 0.005785 | 0.0016021 | 0 | 0.0022338 | 0.0021679 | 0.0067725 | 0.0040825 | 0.0035268 | 0.0030111 | 0.0030984 | 0.0054772 | 0.0016436
| Rel.Std.Dev. |  | 0% | 3.068% | 5.569% | 1.956% | 5.152% | 1.908% | 3.423% | 1.446% | 1.232% | 3.098% | 4.084% | 1.315% | 0% | 1.503% | 1.612% | 2.484% | 2.425% | 1.547% | 2.171% | 1.239% | 3.777% | 1.113%
| PDM3 |  | -13.51% | 12.43% | 13.91% | -7.083% | 1.566% | 8.995% | 6.223% | 2.564% | 6.223% | 10.88% | -5.752% | -18.95% | -6.861% | -1.129% | -10.52% | 81.4% | 11.99% | 51.66% | -7.748% | 66.32% | -3.535% | -1.773%
| Z-Score (Absolute) |  | 1.16 | 0.77 | 0.88 | 0.68 | 0.04 | 0.52 | 0.31 | 0.04 | 0.31 | 0.66 | 0.58 | 1.57 | 0.67 | 0.24 | 0.94 | 5.92 | 0.74 | 3.7 | 0.73 | 4.78 | 0.42 | 0.29
| Analytical results for Tl in OREAS 931b (Certified Value 0.71 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.76841 | 0.69 | 0.69 | 0.7 | 0.69 | 0.7 | 0.77 | 0.62 | 0.72 | < 5 | 0.67 | 0.71 | 0.76 | 0.68 | 0.8 | 0.65 | 0.76 | 0.51683 | 0.72 | 0.6587 |  |  |  | 1
| 1 | 2 | 0.75992 | 0.73 | 0.7 | 0.7 | 0.74 | 0.7 | 0.76 | 0.65 | 0.73 | < 5 | 0.64 | 0.71 | 0.76 | 0.72 | 0.7 | 0.67 | 0.75 | 0.50753 | 0.74 | 0.6703 |  |  |  | 0
| 1 | 3 | 0.76184 | 0.72 | 0.69 | 0.68 | 0.69 | 0.68 | 0.67 | 0.67 | 0.67 | < 5 | 0.69 | 0.74 | 0.74 | 0.71 | 0.7 | 0.68 | 0.77 | 0.54543 | 0.74 | 0.689 |  |  |  | 16
| 1 | 4 | 0.76522 | 0.72 | 0.72 | 0.68 | 0.67 | 0.7 | 0.76 | 0.62 | 0.7 | < 5 | 0.71 | 0.7 | 0.76 | 0.72 | 0.7 | 0.73 | 0.76 | 0.56137 | 0.7 | 0.7193 |  |  |  | 0.70977
| 1 | 5 | 0.75397 | 0.72 | 0.69 | 0.71 | 0.67 | 0.69 | 0.75 | 0.63 | 0.72 | < 5 | 0.69 | 0.71 | 0.76 | 0.72 | 0.7 | 0.76 | 0.76 | 0.55818 | 0.72 | 0.6675 |  |  |  | 57
| 1 | 6 | 0.75261 | 0.69 | 0.67 | 0.7 | 0.68 | 0.69 | 0.72 | 0.69 | 0.7 | < 5 | 0.71 | 0.69 | 0.74 | 0.7 | 0.8 | 0.75 | 0.76 | 0.53622 | 0.71 | 0.6878
| Mean |  | 0.76033 | 0.71167 | 0.69333 | 0.695 | 0.69 | 0.69333 | 0.73833 | 0.64667 | 0.70667 |  | 0.685 | 0.71 | 0.75333 | 0.70833 | 0.73333 | 0.70667 | 0.76 | 0.53759 | 0.72167 | 0.6821
| Median |  | 0.76088 | 0.72 | 0.69 | 0.7 | 0.685 | 0.695 | 0.755 | 0.64 | 0.71 |  | 0.69 | 0.71 | 0.76 | 0.715 | 0.7 | 0.705 | 0.76 | 0.54083 | 0.72 | 0.67905
| Std Dev. |  | 0.0061911 | 0.017224 | 0.01633 | 0.012247 | 0.026077 | 0.008165 | 0.037639 | 0.028752 | 0.021602 |  | 0.026646 | 0.016733 | 0.010328 | 0.016021 | 0.05164 | 0.045898 | 0.0063246 | 0.021851 | 0.016021 | 0.021744
| Rel.Std.Dev. |  | 0.8143% | 2.42% | 2.355% | 1.762% | 3.779% | 1.178% | 5.098% | 4.446% | 3.057% |  | 3.89% | 2.357% | 1.371% | 2.262% | 7.042% | 6.495% | 0.8322% | 4.065% | 2.22% | 3.188%
| PDM3 |  | 7.123% | 0.2672% | -2.316% | -2.081% | -2.785% | -2.316% | 4.024% | -8.891% | -0.4373% |  | -3.49% | 0.03236% | 6.138% | -0.2025% | 3.32% | -0.4373% | 7.077% | -24.26% | 1.676% | -3.898%
| Z-Score (Absolute) |  | 2.17 | 0.2 | 0.54 | 0.47 | 0.67 | 0.54 | 1.28 | 2.43 | 0 | 72.56 | 0.88 | 0.13 | 1.89 | 0.07 | NA | 0 | 2.16 | 6.84 | 0.61 | 0.99
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 931b (Certified Value 0.2 ppm) | [24 empty] | Certified
Tm | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 02: 03; 03: 11; 04: 12; 05: 13; 06: 15; 07: 17; 08: 24 |  |  |  | ppm
|  |  | 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 02: 0.2; 03: 0.1952; 04: 0.18; 05: 0.2; 06: 0.18; 07: 0.3; 08: 0.1585 |  |  |  | 1
| 1 | 2 | 02: 0.2; 03: 0.1915; 04: 0.19; 05: 0.2; 06: 0.19; 07: 0.3; 08: 0.1677 |  |  |  | 6
| 1 | 3 | 02: 0.19; 03: 0.1895; 04: 0.24; 05: 0.2; 06: 0.19; 07: 0.3; 08: 0.1619 |  |  |  | 16
| 1 | 4 | 02: 0.2; 03: 0.1984; 04: 0.21; 05: 0.2; 06: 0.19; 07: 0.3; 08: 0.1707 |  |  |  | 0.19633
| 1 | 5 | 02: 0.2; 03: 0.2014; 04: 0.21; 05: 0.2; 06: 0.19; 07: 0.3; 08: 0.1718 |  |  |  | 58
| 1 | 6 | 02: 0.2; 03: 0.1938; 04: 0.21; 05: 0.2; 06: 0.19; 07: 0.3; 08: 0.1616
| Mean |  | 02: 0.19833; 03: 0.19497; 04: 0.20667; 05: 0.2; 06: 0.18833; 07: 0.3; 08: 0.16537
| Median |  | 02: 0.2; 03: 0.1945; 04: 0.21; 05: 0.2; 06: 0.19; 07: 0.3; 08: 0.1648
| Std Dev. |  | 02: 0.0040825; 03: 0.004393; 04: 0.020656; 05: 3.0405e-17; 06: 0.0040825; 07: 0; 08: 0.0054522
| Rel.Std.Dev. |  | 02: 2.058%; 03: 2.253%; 04: 9.995%; 05: 1.52e-14%; 06: 2.168%; 07: 0%; 08: 3.297%
| PDM3 |  | 02: 1.022%; 03: -0.6927%; 04: 5.267%; 05: 1.871%; 06: -4.071%; 07: 52.81%; 08: -15.77%
| Z-Score (Absolute) |  | 02: 0; 03: 0.27; 04: 0.67; 05: 0.13; 06: 0.81; 07: 8.23; 08: 2.67
| Analytical results for U in OREAS 931b (Certified Value 3.46 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.6997 | 3.39 | 3.47 | 3.4 | 3.6 | 3.5 | 3.5 | 3 | 3.4 | 3.0443 | 3.36 | 3.4 | 3.55 | 3.42 | 3.5 | 3.5 | 3.6 | 3.2523 | 3.53 | 9 | 3.5046 |  |  | 1
| 1 | 2 | 3.6787 | 3.39 | 3.44 | 3.5 | 3.7 | 3.5 | 3.4 | 3.2 | 3.5 | 3.1035 | 3.39 | 3.3 | 3.62 | 3.44 | 3.5 | 3.5 | 3.6 | 3.2641 | 3.73 | 9 | 3.5114 |  |  | 13
| 1 | 3 | 3.6888 | 3.38 | 3.52 | 3.4 | 3.3 | 3.3 | 3.3 | 3.2 | 3.2 | 3.0141 | 3.68 | 3.3 | 3.55 | 3.35 | 3.6 | 3.6 | 3.7 | 3.2342 | 3.69 | 9 | 3.6135 |  |  | 16
| 1 | 4 | 3.6603 | 3.48 | 3.52 | 3.5 | 3.2 | 3.4 | 3.5 | 3.1 | 3.5 | 3.2266 | 3.79 | 3.2 | 3.65 | 3.63 | 3.4 | 3.8 | 3.6 | 3.1765 | 3.73 | 9 | 3.7111 |  |  | 3.4555
| 1 | 5 | 3.6765 | 3.43 | 3.5 | 3.5 | 3.3 | 3.2 | 3.2 | 3.1 | 3.5 | 3.1675 | 3.71 | 3.3 | 3.67 | 3.45 | 3.5 | 3.9 | 3.7 | 3.2144 | 3.69 | 10 | 3.4939 |  |  | 59
| 1 | 6 | 3.6613 | 3.33 | 3.5 | 3.5 | 3.3 | 3.3 | 3.4 | 3.2 | 3.4 | 3.1128 | 3.76 | 3.3 | 3.57 | 3.32 | 3.6 | 3.9 | 3.4 | 3.2034 | 3.63 | 9 | 3.6196
| Mean |  | 3.6775 | 3.4 | 3.4917 | 3.4667 | 3.4 | 3.3667 | 3.3833 | 3.1333 | 3.4167 | 3.1115 | 3.615 | 3.3 | 3.6017 | 3.435 | 3.5167 | 3.7 | 3.6 | 3.2242 | 3.6667 | 9.1667 | 3.5757
| Median |  | 3.6776 | 3.39 | 3.5 | 3.5 | 3.3 | 3.35 | 3.4 | 3.15 | 3.45 | 3.1082 | 3.695 | 3.3 | 3.595 | 3.43 | 3.5 | 3.7 | 3.6 | 3.2243 | 3.69 | 9 | 3.5625
| Std Dev. |  | 0.015388 | 0.050596 | 0.031252 | 0.05164 | 0.2 | 0.12111 | 0.1169 | 0.08165 | 0.1169 | 0.078041 | 0.19003 | 0.063246 | 0.052313 | 0.10858 | 0.075277 | 0.18974 | 0.10954 | 0.032508 | 0.076333 | 0.40825 | 0.086684
| Rel.Std.Dev. |  | 0.4184% | 1.488% | 0.895% | 1.49% | 5.882% | 3.597% | 3.455% | 2.606% | 3.422% | 2.508% | 5.257% | 1.917% | 1.452% | 3.161% | 2.141% | 5.128% | 3.043% | 1.008% | 2.082% | 4.454% | 2.424%
| PDM3 |  | 6.427% | -1.605% | 1.047% | 0.3239% | -1.605% | -2.57% | -2.088% | -9.323% | -1.123% | -9.955% | 4.617% | -4.499% | 4.231% | -0.5926% | 1.771% | 7.076% | 4.182% | -6.694% | 6.112% | 165.3% | 3.479%
| Z-Score (Absolute) |  | 1.12 | 0.25 | 0.2 | 0.08 | 0.25 | 0.41 | 0.33 | 1.56 | 0.17 | 1.67 | 0.81 | 0.74 | 0.74 | 0.08 | 0.32 | 1.23 | 0.73 | 1.12 | 1.06 | NA | 0.61
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 931b (Certified Value 138 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 141 | 146.995 | 144 | 139 | 134 | 135 | 140 | 138 | 131 | 122 | 149.1075 | 149 | 139 | 148 | 135 | 132 | 131 | 141 | 143.7909 | 142 | 165 | 133 | 127.9644 | 1
| 1 | 2 | 141 | 147.0493 | 147 | 138 | 137 | 135 | 139 | 134 | 132 | 130 | 147.492 | 151 | 139 | 147 | 142 | 130 | 130 | 141 | 145.8281 | 142 | 165 | 140 | 128.3661 | 0
| 1 | 3 | 142 | 143.7352 | 140 | 138 | 135 | 133 | 137 | 133 | 133 | 122 | 141.8606 | 155 | 138 | 140 | 137 | 135 | 136 | 143 | 148.7209 | 140 | 163 | 138 | 132.2676 | 16
| 1 | 4 | 141 | 143.9801 | 149 | 141 | 134 | 133 | 136 | 134 | 132 | 128 | 150.5124 | 154 | 138 | 153 | 131 | 131 | 148 | 139 | 145.1262 | 140 | 163 | 141 | 131.3732 | 138.195
| 1 | 5 | 140 | 145.5218 | 138 | 138 | 137 | 133 | 131 | 136 | 131 | 131 | 145.1217 | 161 | 138 | 149 | 134 | 129 | 149 | 142 | 146.9572 | 142 | 165 | 134 | 128.4682 | 60
| 1 | 6 | 140 | 147.1989 | 146 | 136 | 136 | 134 | 134 | 135 | 131 | 130 | 144.4111 | 165 | 138 | 147 | 131 | 132 | 161 | 143 | 149.9573 | 141 | 164 | 137 | 130.9655
| Mean |  | 140.8333 | 145.7467 | 144 | 138.3333 | 135.5 | 133.8333 | 136.1667 | 135 | 131.6667 | 127.1667 | 146.4176 | 155.8333 | 138.3333 | 147.3333 | 135 | 131.5 | 142.5 | 141.5 | 146.7301 | 141.1667 | 164.1667 | 137.1667 | 129.9008
| Median |  | 141 | 146.2584 | 145 | 138 | 135.5 | 133.5 | 136.5 | 134.5 | 131.5 | 129 | 146.3068 | 154.5 | 138 | 147.5 | 134.5 | 131.5 | 142 | 141.5 | 146.3927 | 141.5 | 164.5 | 137.5 | 129.7168
| Std Dev. |  | 0.75277 | 1.5863 | 4.2426 | 1.633 | 1.3784 | 0.98319 | 3.3116 | 1.7889 | 0.8165 | 4.1191 | 3.2134 | 6.08 | 0.5164 | 4.2269 | 4.1473 | 2.0736 | 12.2107 | 1.5166 | 2.3002 | 0.98319 | 0.98319 | 3.1885 | 1.8472
| Rel.Std.Dev. |  | 0.5345% | 1.088% | 2.946% | 1.18% | 1.017% | 0.7346% | 2.432% | 1.325% | 0.6201% | 3.239% | 2.195% | 3.902% | 0.3733% | 2.869% | 3.072% | 1.577% | 8.569% | 1.072% | 1.568% | 0.6965% | 0.5989% | 2.325% | 1.422%
| PDM3 |  | 1.909% | 5.465% | 4.201% | 0.1001% | -1.95% | -3.156% | -1.468% | -2.312% | -4.724% | -7.98% | 5.95% | 12.76% | 0.1001% | 6.613% | -2.312% | -4.845% | 3.115% | 2.392% | 6.176% | 2.15% | 18.79% | -0.7441% | -6.002%
| Z-Score (Absolute) |  | 0.37 | 1.11 | 0.85 | 0 | 0.42 | 0.67 | 0.32 | 0.5 | 1 | 1.67 | 1.21 | 2.62 | 0 | 1.35 | 0.5 | 1.02 | 0.62 | 0.47 | 1.26 | 0.42 | 3.87 | 0.17 | 1.26
| Analytical results for W in OREAS 931b (Certified Value 18.8 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 19 | 18.4446 | 18.8 | 18.8 | 18 | 15.2 | 18.9 | 19.3 | 17.2 | 18.5 | 16.702 | 20.2 | 18.3 | 20 | 18.6 | 17.6 | 19.8 | 19.7 | 15.9543 | 19.8 | 26 | 22 | 18.9823 | 1
| 1 | 2 | 20 | 18.4753 | 19.3 | 18.1 | 18.6 | 15.6 | 18.8 | 19.3 | 17.9 | 19.4 | 17.7529 | 20.2 | 18.4 | 20.4 | 19.2 | 17.5 | 19.3 | 19.2 | 16.179 | 20.4 | 25 | 24 | 19.0153 | 15
| 1 | 3 | 21 | 18.5601 | 19.2 | 18.8 | 17.9 | 18.7 | 17.9 | 19.1 | 17.6 | 18.3 | 17.3829 | 22.1 | 18.5 | 19.9 | 17.8 | 18.6 | 19.2 | 19.5 | 16.2717 | 20.3 | 25 | 26 | 19.5757 | 16
| 1 | 4 | 19 | 18.5255 | 19.5 | 19.1 | 17.4 | 18.9 | 17.8 | 19.6 | 16.9 | 19 | 18.17 | 22.3 | 18 | 20.6 | 19.2 | 17.6 | 18.4 | 19.3 | 15.9409 | 20 | 24 | 23 | 20.199 | 18.7815
| 1 | 5 | 20 | 18.51 | 19.4 | 18.6 | 18.3 | 18.7 | 17.6 | 18.9 | 18.2 | 19.6 |  | 22.6 | 18.4 | 20.8 | 19 | 18.1 | 18.9 | 19.2 | 15.7569 | 20.2 | 26 | 24 | 19.0369 | 61
| 1 | 6 | 19 | 18.4388 | 19.1 | 18.5 | 18.2 | 19.2 | 18 | 19.3 | 17.9 | 18.7 | 18.149 | 22.4 | 17.8 | 20.3 | 18.6 | 17.8 | 18.6 | 19.6 | 16.1096 | 20.5 | 25 | 24 | 19.4003
| Mean |  | 19.6667 | 18.4924 | 19.2167 | 18.65 | 18.0667 | 17.7167 | 18.1667 | 19.25 | 17.6167 | 18.9167 | 17.6314 | 21.6333 | 18.2333 | 20.3333 | 18.7333 | 17.8667 | 19.0333 | 19.4167 | 16.0354 | 20.2 | 25.1667 | 23.8333 | 19.3682
| Median |  | 19.5 | 18.4927 | 19.25 | 18.7 | 18.1 | 18.7 | 17.95 | 19.3 | 17.75 | 18.85 | 17.7529 | 22.2 | 18.35 | 20.35 | 18.8 | 17.7 | 19.05 | 19.4 | 16.0319 | 20.25 | 25 | 24 | 19.2186
| Std Dev. |  | 0.8165 | 0.047858 | 0.24833 | 0.33912 | 0.40825 | 1.8082 | 0.5465 | 0.23452 | 0.48751 | 0.51153 | 0.61203 | 1.1219 | 0.27325 | 0.34448 | 0.53166 | 0.41793 | 0.50859 | 0.2137 | 0.18711 | 0.26077 | 0.75277 | 1.3292 | 0.47276
| Rel.Std.Dev. |  | 4.152% | 0.2588% | 1.292% | 1.818% | 2.26% | 10.21% | 3.008% | 1.218% | 2.767% | 2.704% | 3.471% | 5.186% | 1.499% | 1.694% | 2.838% | 2.339% | 2.672% | 1.101% | 1.167% | 1.291% | 2.991% | 5.577% | 2.441%
| PDM3 |  | 4.713% | -1.539% | 2.317% | -0.7002% | -3.806% | -5.67% | -3.274% | 2.494% | -6.202% | 0.7197% | -6.124% | 15.18% | -2.919% | 8.263% | -0.2565% | -4.871% | 1.341% | 3.382% | -14.62% | 7.553% | 34% | 26.9% | 3.124%
| Z-Score (Absolute) |  | NA | 0.21 | 0.54 | 0.04 | 0.65 | 1.01 | 0.54 | 0.58 | 1.11 | 0.23 | 1.1 | 3.05 | 0.47 | 1.7 | 0.04 | 0.85 | 0.35 | 0.75 | 2.75 | 1.56 | NA | NA | 0.7
Results from laboratories 1, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 931b (Certified Value 13.5 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 1
1
| 1 | 1 | 13 | 12.843 | 12.66 | 20.9 | 15.4 | 13 | 15 | 15.5 | 15.8 | 14.4 | 13.2301 | 8.9 | 15.2 | 12.4 | 11.46 | 13 | 17.6 | 14.6 | 13.7985 | 11.5 | 20 | 13.1 | 10.199 | 1
| 1 | 2 | 13 | 12.6171 | 12.55 | 20.3 | 15.3 | 14.4 | 15.4 | 16 | 16.2 | 15.7 | 12.8602 | 9.5 | 15.2 | 12.7 | 12.71 | 12.6 | 17.5 | 14.8 | 12.6394 | 11.7 | 19 | 12.4 | 10.219 | 16
| 1 | 3 | 13 | 12.7975 | 12.77 | 20.2 | 15 | 13.9 | 13.8 | 15.3 | 16.1 | 14.5 | 12.3743 | 10.1 | 15 | 12.6 | 11.25 | 12.4 | 17.1 | 14.8 | 13.4009 | 11.3 | 19 | 11.9 | 10.2387 | 16
| 1 | 4 | 13 | 12.4852 | 12.67 | 20.8 | 15.4 | 13.6 | 13.8 | 16.1 | 15.6 | 16.3 | 12.6382 | 10 | 15 | 12.9 | 12 | 12.4 | 18.7 | 14 | 11.6978 | 11.7 | 19 | 12 | 10.2888 | 13.5386
| 1 | 5 | 13 | 12.6762 | 12.55 | 20.8 | 15.7 | 13.7 | 13.8 | 15 | 15.8 | 16 | 12.708 | 10.8 | 15.3 | 12.8 | 11.79 | 14.4 | 20.7 | 14.9 | 15.3134 | 12.5 | 20 | 12.4 | 10.089 | 62
| 1 | 6 | 12 | 12.3927 | 12.96 | 20.4 | 15.4 | 13.4 | 13.8 | 15.4 | 15.2 | 15 | 12.6906 | 10.1 | 15.3 | 12.5 | 11.7 | 13.5 | 21.1 | 14.2 | 15.2509 | 11.9 | 20 | 12.7 | 10.4288
| Mean |  | 12.8333 | 12.6353 | 12.6933 | 20.5667 | 15.3667 | 13.6667 | 14.2667 | 15.55 | 15.7833 | 15.3167 | 12.7502 | 9.9 | 15.1667 | 12.65 | 11.8183 | 13.05 | 18.7833 | 14.55 | 13.6835 | 11.7667 | 19.5 | 12.4167 | 10.2439
| Median |  | 13 | 12.6466 | 12.665 | 20.6 | 15.4 | 13.65 | 13.8 | 15.45 | 15.8 | 15.35 | 12.6993 | 10.05 | 15.2 | 12.65 | 11.745 | 12.8 | 18.15 | 14.7 | 13.5997 | 11.7 | 19.5 | 12.4 | 10.2288
| Std Dev. |  | 0.40825 | 0.1748 | 0.15475 | 0.30111 | 0.22509 | 0.47188 | 0.73394 | 0.42308 | 0.36009 | 0.79854 | 0.2834 | 0.64187 | 0.13663 | 0.18708 | 0.50878 | 0.78422 | 1.7279 | 0.36742 | 1.4311 | 0.41312 | 0.54772 | 0.4446 | 0.11212
| Rel.Std.Dev. |  | 3.181% | 1.383% | 1.219% | 1.464% | 1.465% | 3.453% | 5.144% | 2.721% | 2.281% | 5.214% | 2.223% | 6.484% | 0.9008% | 1.479% | 4.305% | 6.009% | 9.199% | 2.525% | 10.46% | 3.511% | 2.809% | 3.581% | 1.094%
| PDM3 |  | -5.21% | -6.672% | -6.244% | 51.91% | 13.5% | 0.9456% | 5.377% | 14.86% | 16.58% | 13.13% | -5.823% | -26.88% | 12.03% | -6.564% | -12.71% | -3.609% | 38.74% | 7.47% | 1.07% | -13.09% | 44.03% | -8.287% | -24.34%
| Z-Score (Absolute) |  | NA | 0.46 | 0.44 | 3.1 | 0.76 | 0 | 0.27 | 0.85 | 0.95 | 0.74 | 0.41 | 1.69 | 0.67 | 0.46 | 0.83 | 0.28 | 2.3 | 0.4 | 0.01 | 0.85 | NA | 0.56 | 1.54
Results from laboratories 1 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 931b (Certified Value 1.38 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 03; 02: 10; 03: 11; 04: 12; 05: 13; 06: 15; 07: 17; 08: 20; 09: 22; 10: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*OES; 10: 4A*OES/MS | 2
2
| 1 | 1 | 01: 1.44; 02: 1.5; 03: 1.2628; 04: 1.3; 05: 1.6; 06: 1.31; 07: 1.8; 08: 1.4; 09: 2; 10: 1.157 | 1
| 1 | 2 | 01: 1.39; 02: 1.6; 03: 1.2424; 04: 1.2; 05: 1.6; 06: 1.34; 07: 1.8; 08: 1.4; 09: 2; 10: 1.1383 | 0
| 1 | 3 | 01: 1.32; 02: 1.5; 03: 1.2238; 04: 1.7; 05: 1.6; 06: 1.28; 07: 1.8; 08: 1.3; 09: 2; 10: 1.1583 | 16
| 1 | 4 | 01: 1.39; 02: 1.6; 03: 1.2641; 04: 1.4; 05: 1.5; 06: 1.34; 07: 2; 08: 1.4; 09: 2; 10: 1.1884 | 1.3773
| 1 | 5 | 01: 1.37; 02: 1.6; 03: 1.2942; 04: 1.4; 05: 1.5; 06: 1.33; 07: 2.1; 08: 1.5; 09: 2; 10: 1.1106 | 63
| 1 | 6 | 01: 1.31; 02: 1.5; 03: 1.2623; 04: 1.5; 05: 1.5; 06: 1.36; 07: 2.1; 08: 1.4; 09: 2; 10: 1.1286
| Mean |  | 01: 1.37; 02: 1.55; 03: 1.2583; 04: 1.4167; 05: 1.55; 06: 1.3267; 07: 1.9333; 08: 1.4; 09: 2; 10: 1.1469
| Median |  | 01: 1.38; 02: 1.55; 03: 1.2626; 04: 1.4; 05: 1.55; 06: 1.335; 07: 1.9; 08: 1.4; 09: 2; 10: 1.1477
| Std Dev. |  | 01: 0.04858; 02: 0.054772; 03: 0.023672; 04: 0.17224; 05: 0.054772; 06: 0.028048; 07: 0.15055; 08: 0.063246; 09: 0; 10: 0.027132
| Rel.Std.Dev. |  | 01: 3.546%; 02: 3.534%; 03: 1.881%; 04: 12.16%; 05: 3.534%; 06: 2.114%; 07: 7.787%; 08: 4.518%; 09: 0%; 10: 2.366%
| PDM3 |  | 01: -0.5306%; 02: 12.54%; 03: -8.643%; 04: 2.858%; 05: 12.54%; 06: -3.677%; 07: 40.37%; 08: 1.648%; 09: 45.21%; 10: -16.73%
| Z-Score (Absolute) |  | 01: 0.14; 02: 0.71; 03: 0.67; 04: 0.08; 05: 0.71; 06: 0.35; 07: 2.54; 08: 0; 09: NA; 10: 1.2
Results from laboratory 22 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Zn in OREAS 931b (Certified Value 287 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 0
0
| 1 | 1 | 293 | 300.7485 | 290 | 291 | 281 | 289 | 291 | 280 | 286 | 272 | 277.4612 | 297 | 289.7 | 295 | 280 | 290 | 305 | 293 | 280.5508 | 282 | 358 | 293 | 272.0338 | 1
| 1 | 2 | 289 | 301.4281 | 290 | 290 | 285 | 284 | 289 | 276 | 292 | 281 | 273.5836 | 305 | 290.3 | 294 | 288 | 289 | 304 | 291 | 294.8199 | 281 | 361 | 295 | 272.7605 | 0
| 1 | 3 | 290 | 299.3976 | 288 | 289 | 283 | 267 | 285 | 272 | 289 | 263 | 260.6107 | 302 | 291.2 | 281 | 273 | 289 | 301 | 296 | 286.1619 | 276 | 363 | 292 | 280.4685 | 16
| 1 | 4 | 290 | 301.9458 | 295 | 287 | 278 | 269 | 283 | 273 | 286 | 277 | 285.7904 | 305 | 291 | 299 | 293 | 280 | 313 | 293 | 286.4751 | 276 | 367 | 290 | 276.4731 | 286.9808
| 1 | 5 | 290 | 302.6742 | 286 | 286 | 284 | 270 | 274 | 278 | 287 | 283 | 269.6014 | 320 | 292.1 | 298 | 284 | 286 | 312 | 294 | 279.9271 | 281 | 357 | 289 | 274.3609 | 64
| 1 | 6 | 291 | 299.3498 | 285 | 287 | 284 | 269 | 279 | 275 | 296 | 264 | 268.1514 | 331 | 288.5 | 289 | 281 | 300 | 316 | 298 | 289.1769 | 280 | 368 | 291 | 275.1142
| Mean |  | 290.5 | 300.924 | 289 | 288.3333 | 282.5 | 274.6667 | 283.5 | 275.6667 | 289.3333 | 273.3333 | 272.5331 | 310 | 290.4667 | 292.6667 | 283.1667 | 289 | 308.5 | 294.1667 | 286.1853 | 279.3333 | 362.3333 | 291.6667 | 275.2018
| Median |  | 290 | 301.0883 | 289 | 288 | 283.5 | 269.5 | 284 | 275.5 | 288 | 274.5 | 271.5925 | 305 | 290.65 | 294.5 | 282.5 | 289 | 308.5 | 293.5 | 286.3185 | 280.5 | 362 | 291.5 | 274.7375
| Std Dev. |  | 1.3784 | 1.3565 | 3.5777 | 1.9664 | 2.5884 | 9.3524 | 6.3166 | 3.0111 | 3.9833 | 8.501 | 8.6149 | 12.8374 | 1.2628 | 6.7132 | 6.9113 | 6.5115 | 5.9582 | 2.4833 | 5.5588 | 2.6583 | 4.5461 | 2.1602 | 3.0353
| Rel.Std.Dev. |  | 0.4745% | 0.4508% | 1.238% | 0.682% | 0.9163% | 3.405% | 2.228% | 1.092% | 1.377% | 3.11% | 3.161% | 4.141% | 0.4347% | 2.294% | 2.441% | 2.253% | 1.931% | 0.8442% | 1.942% | 0.9517% | 1.255% | 0.7407% | 1.103%
| PDM3 |  | 1.226% | 4.859% | 0.7036% | 0.4713% | -1.561% | -4.291% | -1.213% | -3.942% | 0.8198% | -4.756% | -5.034% | 8.021% | 1.215% | 1.981% | -1.329% | 0.7036% | 7.498% | 2.504% | -0.2772% | -2.665% | 26.26% | 1.633% | -4.104%
| Z-Score (Absolute) |  | 0.17 | 1.38 | 0 | 0.08 | 0.75 | 1.66 | 0.64 | 1.54 | 0.04 | 1.81 | 1.9 | 2.43 | 0.17 | 0.42 | 0.67 | 0 | 2.25 | 0.6 | 0.33 | 1.12 | 8.48 | 0.31 | 1.6
| Analytical results for Zr in OREAS 931b (Certified Value 62 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 0
1
| 1 | 1 | 62 | 61.3405 | 63.7 | 66.9 | 62.4 | 55.6 | 66.5 | 66 | 65 | 58.4 | 60.0888 | 63.5 | 66.6 | 65.5 | 58.2 | 58.7 | 64 | 63.5 | 83.0768 | 58.5 | 71 | 58.1 | 50.3363 | 1
| 1 | 2 | 62 | 59.9102 | 63.7 | 67.1 | 62.3 | 57.5 | 65.1 | 66.8 | 66.1 | 61.4 | 63.9521 | 63.5 | 66.9 | 63.9 | 62.3 | 59.8 | 66 | 65.3 | 80.201 | 58.4 | 71 | 52.7 | 50.9698 | 19
| 1 | 3 | 62 | 61.4134 | 65.1 | 66.1 | 62.6 | 63 | 62.1 | 66 | 67.1 | 57 | 63.6359 | 66.1 | 66.9 | 60.4 | 57.9 | 58.7 | 66 | 65.9 | 82.7858 | 57.5 | 72 | 51.3 | 52.9997 | 16
| 1 | 4 | 63 | 60.5554 | 65.7 | 69.2 | 60.9 | 60.8 | 63 | 66.7 | 66.2 | 59.1 | 65.09 | 67.2 | 66.8 | 65.5 | 62.3 | 56.4 | 67 | 62.5 | 79.5014 | 57.8 | 72 | 53.6 | 52.8003 | 62.1344
| 1 | 5 | 62 | 60.5512 | 64.1 | 67.9 | 61 | 62.2 | 61.2 | 64.7 | 66.6 | 60 | 64.85 | 71.7 | 66.7 | 65.8 | 61 | 60 | 72 | 64.9 | 81.447 | 58.5 | 71 | 60 | 50.8577 | 65
| 1 | 6 | 61 | 59.7198 | 64.6 | 66.8 | 61.6 | 62.3 | 61.3 | 66.3 | 66.1 | 58.9 | 66.0463 | 74 | 66 | 63.5 | 59.3 | 61 | 78 | 64.7 | 83.1269 | 57.7 | 72 | 53.7 | 52.1836
| Mean |  | 62 | 60.5817 | 64.4833 | 67.3333 | 61.8 | 60.2333 | 63.2 | 66.0833 | 66.1833 | 59.1333 | 63.9438 | 67.6667 | 66.65 | 64.1 | 60.1667 | 59.1 | 68.8333 | 64.4667 | 81.6898 | 58.0667 | 71.5 | 54.9 | 51.6912
| Median |  | 62 | 60.5533 | 64.35 | 67 | 61.95 | 61.5 | 62.55 | 66.15 | 66.15 | 59 | 64.401 | 66.65 | 66.75 | 64.7 | 60.15 | 59.25 | 66.5 | 64.8 | 82.1164 | 58.1 | 71.5 | 53.65 | 51.5767
| Std Dev. |  | 0.63246 | 0.70185 | 0.80602 | 1.082 | 0.74027 | 3.0018 | 2.1615 | 0.75741 | 0.6969 | 1.4855 | 2.0742 | 4.3307 | 0.33912 | 2.0445 | 1.9775 | 1.5824 | 5.2313 | 1.2485 | 1.5656 | 0.45019 | 0.54772 | 3.3817 | 1.1168
| Rel.Std.Dev. |  | 1.02% | 1.159% | 1.25% | 1.607% | 1.198% | 4.984% | 3.42% | 1.146% | 1.053% | 2.512% | 3.244% | 6.4% | 0.5088% | 3.19% | 3.287% | 2.678% | 7.6% | 1.937% | 1.916% | 0.7753% | 0.766% | 6.16% | 2.16%
| PDM3 |  | -0.2164% | -2.499% | 3.78% | 8.367% | -0.5383% | -3.06% | 1.715% | 6.355% | 6.516% | -4.83% | 2.912% | 8.904% | 7.267% | 3.163% | -3.167% | -4.884% | 10.78% | 3.754% | 31.47% | -6.547% | 15.07% | -11.64% | -16.81%
| Z-Score (Absolute) |  | NA | 0.63 | 0.19 | 0.79 | 0.37 | 0.7 | 0.08 | 0.53 | 0.55 | 0.93 | 0.08 | 0.86 | 0.65 | 0.11 | 0.72 | 0.94 | NA | 0.19 | 3.81 | 1.16 | NA | 1.82 | 2.5
Results from laboratories 1, 17 and 22 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C362: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C473: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C491: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C619: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C656: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C675: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C693: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C711: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C729: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C747: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C765: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C783: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C802: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C934: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1083: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1102: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1121: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 931b (Certified Value 17.1 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | 15.1 | 17.1691 | 16.48 | 14.7 | 18.7 | 16.3 | 17.5 | 16.85 | 16 | 18.39 | 15.97 | > 10 | 15.11 | 18.1 | 16.69 | 15.8 | 16.3 | 17.22 | 15.9109 | 18.93 | 19.4 | 17 | 17.2766 | 1
| 1 | 2 | 14.9 | 17.0604 | 16.4 | 14.9 | 18.35 | 16.3 | 17.5 | 16.75 | 17.45 | 18.3 | 15.31 | > 10 | 16.83 | 18.3 | 16.68 | 16.6 | 16 | 18.01 | 14.5156 | 18.27 | 19.9 | 17 | 17.6914 | 19
| 1 | 3 | 14.9 | 17.3034 | 16.8 | 15 | 18.15 | 16.65 | 18 | 16.65 | 16.75 | 18.48 | 16.1 | > 10 | 17.541 | 18.4 | 16.86 | 15.8 | 16.4 | 18.3 | 20.0286 | 18.68 | 19.6 | 17 | 18.6372 | 16
| 1 | 4 | 14.9 | 17.2953 | 17.06 | 14.8 | 18.85 | 16.45 | 17.45 | 16.65 | 16.75 | 18.07 | 15.67 | > 10 | 15.841 | 18.8 | 16.78 | 16.1 | 15.7 | 17.87 | 18.5421 | 18.76 | 19.9 | 17 | 18.3519 | 17.1053
| 1 | 5 | 15 | 17.0209 | 16.79 | 15 | 18.1 | 16.6 | 17.7 | 16.65 | 17 | 18.13 | 15.6 | > 10 | 16.696 | 18.7 | 16.17 | 16.3 | 16.3 | 17.31 | 18.6653 | 18.57 | 19.1 | 18 | 18.6247 | 67
| 1 | 6 | 15 | 17.0192 | 16.88 | 14.9 | 18.3 | 16.9 | 17.75 | 16.2 | 16.1 | 18.35 | 15.66 | > 10 | 16.48 | 18.3 | 16.15 | 16.1 | 16 | 17.6 | 19.5883 | 18.23 | 19.5 | 17 | 18.6955
| Mean |  | 14.9667 | 17.1447 | 16.735 | 14.8833 | 18.4083 | 16.5333 | 17.65 | 16.625 | 16.675 | 18.2867 | 15.7183 |  | 16.4163 | 18.4333 | 16.555 | 16.1167 | 16.1167 | 17.7183 | 17.8751 | 18.5733 | 19.5667 | 17.1667 | 18.2129
| Median |  | 14.95 | 17.1148 | 16.795 | 14.9 | 18.325 | 16.525 | 17.6 | 16.65 | 16.75 | 18.325 | 15.665 |  | 16.588 | 18.35 | 16.685 | 16.1 | 16.15 | 17.735 | 18.6037 | 18.625 | 19.55 | 17 | 18.4883
| Std Dev. |  | 0.08165 | 0.13164 | 0.24946 | 0.1169 | 0.30235 | 0.23166 | 0.20976 | 0.22305 | 0.54841 | 0.15731 | 0.28124 |  | 0.8433 | 0.26583 | 0.31297 | 0.30605 | 0.26394 | 0.41845 | 2.1811 | 0.27689 | 0.30768 | 0.40825 | 0.59168
| Rel.Std.Dev. |  | 0.5455% | 0.7678% | 1.491% | 0.7855% | 1.642% | 1.401% | 1.188% | 1.342% | 3.289% | 0.8602% | 1.789% |  | 5.137% | 1.442% | 1.89% | 1.899% | 1.638% | 2.362% | 12.2% | 1.491% | 1.572% | 2.378% | 3.249%
| PDM3 |  | -12.5% | 0.2306% | -2.165% | -12.99% | 7.618% | -3.344% | 3.185% | -2.808% | -2.515% | 6.907% | -8.108% |  | -4.028% | 7.764% | -3.217% | -5.78% | -5.78% | 3.584% | 4.501% | 8.583% | 14.39% | 0.3589% | 6.475%
| Z-Score (Absolute) |  | 1.21 | 0.28 | 0 | 1.27 | 1.15 | 0.14 | 0.63 | 0.08 | 0.04 | 1.06 | 0.7 | NA | 0.22 | 1.16 | 0.12 | 0.42 | 0.42 | 0.67 | 0.78 | 1.26 | 1.94 | NA | 1.01
Results from laboratory 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Al in OREAS 931b (Certified Value 0.727 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  | 3
| 1 | 1 | 1 | 0.93895 | 0.703 | 0.824 | 0.77 | 0.7 | 0.74 | 0.69 | 0.7 | 0.78 | 0.72403 | 0.782 | 0.82 | 0.58 | 1.0528 | 0.71 | 0.97 | 0.72 | 0.7704 | 0.71 | 0.651 | 0.725 |  | 1
| 1 | 2 | 1 | 0.92978 | 0.704 | 0.801 | 0.79 | 0.71 | 0.74 | 0.69 | 0.69 | 0.76 | 0.71586 | 0.799 | 0.97 | 0.58 | 1.0638 | 0.74 | 0.94 | 0.74 | 0.75828 | 0.71 | 0.677 | 0.722 |  | 0
| 1 | 3 | 1.01 | 0.94063 | 0.705 | 0.864 | 0.76 | 0.72 | 0.74 | 0.7 | 0.7 | 0.76 | 0.72391 | 0.797 | 0.99 | 0.59 | 0.9834 | 0.7 | 0.87 | 0.75 | 0.78168 | 0.72 | 0.63 | 0.716 |  | 16
| 1 | 4 | 0.99 | 0.86103 | 0.702 | 0.897 | 0.77 | 0.71 | 0.75 | 0.71 | 0.69 | 0.77 | 0.737 | 0.793 | 0.86 | 0.63 | 1.0462 | 0.7 | 0.94 | 0.73 | 0.80928 | 0.75 | 0.614 | 0.731 |  | 0.72738
| 1 | 5 | 0.99 | 0.8931 | 0.703 | 0.893 | 0.76 | 0.72 | 0.74 | 0.7 | 0.69 | 0.79 | 0.72202 | 0.759 | 0.92 | 0.61 | 1.065 | 0.73 | 0.9 | 0.73 | 0.71448 | 0.73 | 0.656 | 0.737 |  | 68
| 1 | 6 | 0.99 | 0.85883 | 0.699 | 0.873 | 0.76 | 0.73 | 0.71 | 0.68 | 0.66 | 0.79 | 0.72549 | 0.817 | 0.95 | 0.6 | 1.0312 | 0.72 | 0.95 | 0.74 | 0.72216 | 0.73 | 0.614 | 0.714
| Mean |  | 0.99667 | 0.90372 | 0.70267 | 0.85867 | 0.76833 | 0.715 | 0.73667 | 0.695 | 0.68833 | 0.775 | 0.72472 | 0.79117 | 0.91833 | 0.59833 | 1.0404 | 0.71667 | 0.92833 | 0.735 | 0.75938 | 0.725 | 0.64033 | 0.72417
| Median |  | 0.995 | 0.91144 | 0.703 | 0.8685 | 0.765 | 0.715 | 0.74 | 0.695 | 0.69 | 0.775 | 0.72397 | 0.795 | 0.935 | 0.595 | 1.0495 | 0.715 | 0.94 | 0.735 | 0.76434 | 0.725 | 0.6405 | 0.7235
| Std Dev. |  | 0.008165 | 0.038033 | 0.0020656 | 0.038495 | 0.01169 | 0.010488 | 0.013663 | 0.010488 | 0.01472 | 0.013784 | 0.0069033 | 0.019416 | 0.066156 | 0.019408 | 0.030574 | 0.01633 | 0.03656 | 0.010488 | 0.036087 | 0.015166 | 0.025288 | 0.0087958
| Rel.Std.Dev. |  | 0.8192% | 4.209% | 0.294% | 4.483% | 1.522% | 1.467% | 1.855% | 1.509% | 2.138% | 1.779% | 0.9525% | 2.454% | 7.204% | 3.244% | 2.939% | 2.279% | 3.938% | 1.427% | 4.752% | 2.092% | 3.949% | 1.215%
| PDM3 |  | 37.02% | 24.24% | -3.398% | 18.05% | 5.63% | -1.702% | 1.277% | -4.452% | -5.368% | 6.547% | -0.366% | 8.769% | 26.25% | -17.74% | 43.03% | -1.473% | 27.63% | 1.047% | 4.399% | -0.3274% | -11.97% | -0.4419%
| Z-Score (Absolute) |  | 4.4 | 2.83 | 0.56 | 2.07 | 0.55 | 0.35 | 0.01 | 0.69 | 0.8 | 0.66 | 0.19 | 0.93 | 3.08 | 2.32 | 5.13 | 0.32 | 3.25 | 0.01 | 0.4 | 0.18 | 1.61 | 0.2
| Analytical results for As in OREAS 931b (Certified Value 97 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | 0
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 1
| 1 | 1 | 104 | 92.5991 | 96.37 | 91 | 103.5 | 95.5 | 100 | 94.7 | 88.8 | 103 | 93.957 | 95 | 90.2 | 95 | 102 | 88.9 | 96.5 | 98.6 | 104.4139 | 109 | 89 | 99 | 96.0294 | 1
| 1 | 2 | 105 | 94.5588 | 94.36 | 90 | 102 | 94.7 | 100 | 95.7 | 94.8 | 103 | 86.566 | 92 | 101.2 | 94.8 | 101 | 92 | 94.3 | 103.4 | 100.295 | 104 | 88 | 99 | 99.8411 | 0
| 1 | 3 | 105 | 95.2862 | 96.06 | 90 | 102 | 96.8 | 101 | 93.2 | 91.1 | 104 | 92.115 | 98 | 107.2 | 94.7 | 105 | 88.5 | 94.6 | 103.2 | 112.0236 | 111 | 79 | 89 | 88.6593 | 16
| 1 | 4 | 105 | 94.5938 | 97.5 | 88 | 103.5 | 96.8 | 99.7 | 94.4 | 92.1 | 102 | 94.513 | 97 | 93.1 | 95.9 | 103 | 87.9 | 91.5 | 100.8 | 105.3773 | 110 | 88 | 92 | 103.5116 | 97.1731
| 1 | 5 | 105 | 92.0339 | 94.71 | 91 | 102.5 | 95.9 | 99.9 | 92.6 | 92.4 | 104 | 92.405 | 92 | 99.2 | 96 | 101 | 89.6 | 95.6 | 97.9 | 98.7204 | 110 | 85 | 94 | 102.7406 | 69
| 1 | 6 | 105 | 93.649 | 97.34 | 89 | 101.5 | 97.7 | 102.5 | 93.2 | 87.4 | 104 | 91.997 | 98 | 100.6 | 92.6 | 100 | 90 | 93.6 | 99 | 107.6986 | 108 | 86 | 90 | 105.4307
| Mean |  | 104.8333 | 93.7868 | 96.0567 | 89.8333 | 102.5 | 96.2333 | 100.5167 | 93.9667 | 91.1 | 103.3333 | 91.9255 | 95.3333 | 98.5833 | 94.8333 | 102 | 89.4833 | 94.35 | 100.4833 | 104.7548 | 108.6667 | 85.8333 | 93.8333 | 99.3688
| Median |  | 105 | 94.1039 | 96.215 | 90 | 102.25 | 96.35 | 100 | 93.8 | 91.6 | 103.5 | 92.26 | 96 | 99.9 | 94.9 | 101.5 | 89.25 | 94.45 | 99.9 | 104.8956 | 109.5 | 87 | 93 | 101.2909
| Std Dev. |  | 0.40825 | 1.2647 | 1.3054 | 1.169 | 0.83666 | 1.0764 | 1.0722 | 1.1639 | 2.6593 | 0.8165 | 2.822 | 2.8048 | 6.0973 | 1.2275 | 1.7889 | 1.4442 | 1.7283 | 2.3836 | 4.8643 | 2.5033 | 3.656 | 4.3551 | 6.1862
| Rel.Std.Dev. |  | 0.3894% | 1.348% | 1.359% | 1.301% | 0.8163% | 1.119% | 1.067% | 1.239% | 2.919% | 0.7902% | 3.07% | 2.942% | 6.185% | 1.294% | 1.754% | 1.614% | 1.832% | 2.372% | 4.644% | 2.304% | 4.259% | 4.641% | 6.226%
| PDM3 |  | 7.883% | -3.485% | -1.149% | -7.553% | 5.482% | -0.9671% | 3.441% | -3.3% | -6.25% | 6.339% | -5.4% | -1.893% | 1.451% | -2.408% | 4.967% | -7.913% | -2.905% | 3.407% | 7.802% | 11.83% | -11.67% | -3.437% | 2.26%
| Z-Score (Absolute) |  | 1.34 | 0.35 | 0 | 0.95 | 0.98 | 0.03 | 0.68 | 0.32 | 0.76 | 1.11 | 0.63 | 0.11 | 0.38 | 0.19 | 0.91 | 1 | 0.26 | 0.67 | 1.32 | 1.92 | 1.56 | 0.34 | 0.5
| Analytical results for Au in OREAS 931b (Certified Value 0.018 ppm) | [24 empty] | Certified
Au | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 13 | 14 | 15 | 16 | 17 | 18 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.001 | 0.0166 | 0.02 | 0.02 | < 0.02 | 0.02 | 0.02 | 0.021 | 0.0275 | 0.033 | 0.0149 | < 0.0005 | 0.0172 | [10 empty] | 1
| 1 | 2 | < 0.001 | 0.0167 | 0.02 | 0.02 | < 0.02 | 0.02 | 0.02 | 0.0105 | 0.0291 | 0.032 | 0.0153 | < 0.0005 | 0.0144 | [10 empty] | 0
| 1 | 3 | < 0.001 | 0.0174 | 0.02 | 0.02 | < 0.02 | < 0.02 | 0.02 | 0.0105 | 0.0286 | 0.034 | 0.0139 | < 0.0005 | 0.0165 | [10 empty] | 16
| 1 | 4 | < 0.001 | 0.0172 | 0.02 | 0.02 | < 0.02 | < 0.02 | 0.02 | 0.0105 | 0.0318 | 0.026 | 0.015 | < 0.0005 | 0.0178 | [10 empty] | 0.017744
| 1 | 5 | < 0.001 | 0.0171 | 0.02 | < 0.02 | 0.02 | 0.02 | 0.02 | 0.0173 | 0.0325 | 0.031 | 0.0133 | < 0.0005 | 0.0168 | [10 empty] | 70
| 1 | 6 | < 0.001 | 0.0176 | 0.02 | < 0.02 | 0.02 | 0.02 | 0.02 | 0.0105 | 0.0306 | 0.03 | 0.014 | < 0.0005 | 0.0153
| Mean |  |  | 0.0171 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.013383 | 0.030017 | 0.031 | 0.0144 |  | 0.016333
| Median |  |  | 0.01715 | 0.02 | 0.02 | 0.02 | 0.02 | 0.02 | 0.0105 | 0.02985 | 0.0315 | 0.01445 |  | 0.01665
| Std Dev. |  |  | 0.00038987 | 0 | 0 | 0 | 0 | 0 | 0.0046175 | 0.0019426 | 0.0028284 | 0.00077974 |  | 0.0012612
| Rel.Std.Dev. |  |  | 2.28% | 0% | 0% | 0% | 0% | 0% | 34.5% | 6.472% | 9.124% | 5.415% |  | 7.722%
| PDM3 |  |  | -3.628% | 12.72% | 12.72% | 12.72% | 12.72% | 12.72% | -24.57% | 69.17% | 74.71% | -18.84% |  | -7.949%
| Z-Score (Absolute) |  | 3.27 | 0.09 | 0.67 | 0 | 0.67 | 0 | 0.67 | 0.66 | 2.7 | 2.9 | 0.46 | 3.32 | 0.07
| Analytical results for B in OREAS 931b (Certified Value < 10 ppm) | [24 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 13 | 14 | 15 | 16 | 17 | 18 | [8 empty] | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 1
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | [8 empty] | 1
| 1 | 1 | 21 | < 10 | < 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | < 1 | < 10 | < 20 | 3 | 17 | [8 empty] | 1
| 1 | 2 | 49 | < 10 | < 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | < 1 | < 10 | < 20 | 3 | 15 | [8 empty] | 1
| 1 | 3 | 50 | < 10 | < 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | < 1 | < 10 | < 20 | 3 | 17 | [8 empty] | 16
| 1 | 4 | 51 | < 10 | < 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | < 1 | < 10 | < 20 | 3 | 14 | [8 empty] | < 10
| 1 | 5 | 49 | < 10 | < 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 1.8 | < 10 | < 20 | 3 | 15 | [8 empty] | 71
| 1 | 6 | 49 | < 10 | < 5 | < 10 | < 10 | < 10 | < 10 | < 10 | < 10 | < 20 | 1.5 | < 10 | < 20 | 2 | 18
| Mean |  | 01: 44.8333; 14: 1.65; 17: 2.8333; 18: 16
| Median |  | 01: 49; 14: 1.65; 17: 3; 18: 16
| Std Dev. |  | 01: 11.7033; 14: 0.21213; 17: 0.40825; 18: 1.5492
| Rel.Std.Dev. |  | 01: 26.1%; 14: 12.86%; 17: 14.41%; 18: 9.682%
PDM3
Z-Score (Absolute) | bar: 01 to 18
| Analytical results for Ba in OREAS 931b (Indicative Value 92 ppm) | [24 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 0
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 0
| 1 | 1 | 25 | 118.8903 | 61.94 | 183 | 40 | 90 | 80 | 80 | 80 | 102 | 126.979 | 123 | 30.7 | 105 | 210 | 122 | 6.8 | 73 | 156.0999 | 83 | 10 | 138 |  | 1
| 1 | 2 | 25 | 126.1047 | 66.26 | 190 | 40 | 90 | 100 | 80 | 80 | 95 | 126.033 | 126 | 26 | 94 | 214 | 85 | 6.5 | 77 | 144.7892 | 97 | 9 | 125 |  | 3
| 1 | 3 | 26 | 126.739 | 64.68 | 182 | 50 | 100 | 50 | 80 | 80 | 90 | 126.976 | 125 | 17.3 | 104 | 195 | 100 | 7.6 | 65 | 148.6482 | 102 | 9 | 134 |  | 16
| 1 | 4 | 25 | 142.8687 | 62.33 | 183 | 40 | 100 | 60 | 90 | 90 | 107 | 129.933 | 125 | 32.4 | 107 | 212 | 77 | 6 | 70 | 161.0134 | 81 | 9 | 139 |  | 91.9789
| 1 | 5 | 24 | 124.3735 | 62.38 | 184 | 40 | 110 | 90 | 80 | 80 | 102 | 125.81 | 120 | 22.9 | 103 | 204 | 71 | 7 | 61 | 137.908 | 85 | 11 | 139 |  | 9
| 1 | 6 | 25 | 142.5422 | 64.06 | 180 | 50 | 80 | 110 | 80 | 80 | 103 | 127.938 | 131 | 21.5 | 104 | 201 | 88 | 6.1 | 77 | 136.1126 | 101 | 10 | 134
| Mean |  | 25 | 130.2531 | 63.6083 | 183.6667 | 43.3333 | 95 | 81.6667 | 81.6667 | 81.6667 | 99.8333 | 127.2782 | 125 | 25.1333 | 102.8333 | 206 | 90.5 | 6.6667 | 70.5 | 147.4285 | 91.5 | 9.6667 | 134.8333
| Median |  | 25 | 126.4218 | 63.22 | 183 | 40 | 95 | 85 | 80 | 80 | 102 | 126.9775 | 125 | 24.45 | 104 | 207 | 86.5 | 6.65 | 71.5 | 146.7187 | 91 | 9.5 | 136
| Std Dev. |  | 0.63246 | 10.0341 | 1.6917 | 3.3862 | 5.164 | 10.4881 | 23.1661 | 4.0825 | 4.0825 | 6.1779 | 1.508 | 3.6332 | 5.7288 | 4.535 | 7.2938 | 18.3385 | 0.59889 | 6.5038 | 9.868 | 9.5446 | 0.8165 | 5.3448
| Rel.Std.Dev. |  | 2.53% | 7.704% | 2.66% | 1.844% | 11.92% | 11.04% | 28.37% | 4.999% | 4.999% | 6.188% | 1.185% | 2.907% | 22.79% | 4.41% | 3.541% | 20.26% | 8.983% | 9.225% | 6.693% | 10.43% | 8.447% | 3.964%
| PDM3 |  | -72.82% | 41.61% | -30.84% | 99.68% | -52.89% | 3.285% | -11.21% | -11.21% | -11.21% | 8.539% | 38.38% | 35.9% | -72.67% | 11.8% | 124% | -1.608% | -92.75% | -23.35% | 60.29% | -0.5206% | -89.49% | 46.59%
| Z-Score (Absolute) |  | 1.27 | 0.75 | 0.53 | 1.78 | 0.91 | 0.08 | 0.18 | 0.18 | 0.18 | 0.17 | 0.7 | 0.65 | 1.26 | 0.23 | 2.21 | 0.01 | 1.62 | 0.39 | 1.08 | 0.01 | 1.56 | 0.84
| Analytical results for Be in OREAS 931b (Certified Value 0.44 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 23 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  |  |  | 2
| 1 | 1 | < 5 | < 0.5 | 0.49 | 0.54 | 0.43 | 0.4 | 0.41 | 0.39 | 0.38 | 0.6 | < 5 | 0.47 | 0.26 | 0.6 | 0.5 | 0.41 | 0.73402 | 0.5 | < 0.5 | < 0.5 |  |  |  | 1
| 1 | 2 | < 5 | < 0.5 | 0.51 | 0.51 | 0.43 | 0.41 | 0.41 | 0.39 | 0.38 | 0.5 | < 5 | 0.42 | 0.24 | 0.58 | 0.5 | 0.41 | 0.6396 | 0.5 | < 0.5 | < 0.5 |  |  |  | 0
| 1 | 3 | < 5 | < 0.5 | 0.5 | 0.52 | 0.42 | 0.42 | 0.41 | 0.4 | 0.38 | 0.5 | < 5 | 0.46 | 0.25 | 0.61 | 0.4 | 0.4 | 0.73261 | 0.5 | < 0.5 | < 0.5 |  |  |  | 16
| 1 | 4 | < 5 | < 0.5 | 0.52 | 0.54 | 0.43 | 0.41 | 0.41 | 0.4 | 0.38 | 0.4 | < 5 | 0.45 | 0.29 | 0.6 | 0.4 | 0.4 | 0.75921 | 0.5 | < 0.5 | < 0.5 |  |  |  | 0.43556
| 1 | 5 | < 5 | < 0.5 | 0.52 | 0.51 | 0.42 | 0.41 | 0.41 | 0.4 | 0.38 | 0.5 | < 5 | 0.39 | 0.23 | 0.52 | 0.5 | 0.37 | 0.71364 | 0.5 | < 0.5 | < 0.5 |  |  |  | 72
| 1 | 6 | < 5 | < 0.5 | 0.53 | 0.5 | 0.43 | 0.42 | 0.41 | 0.39 | 0.36 | 0.5 | < 5 | 0.48 | 0.25 | 0.57 | 0.5 | 0.39 | 0.72147 | 0.5 | < 0.5 | < 0.5
| Mean |  |  |  | 0.51167 | 0.52 | 0.42667 | 0.41167 | 0.41 | 0.395 | 0.37667 | 0.5 |  | 0.445 | 0.25333 | 0.58 | 0.46667 | 0.39667 | 0.71676 | 0.5
| Median |  |  |  | 0.515 | 0.515 | 0.43 | 0.41 | 0.41 | 0.395 | 0.38 | 0.5 |  | 0.455 | 0.25 | 0.59 | 0.5 | 0.4 | 0.72704 | 0.5
| Std Dev. |  |  |  | 0.01472 | 0.016733 | 0.005164 | 0.0075277 | 0 | 0.0054772 | 0.008165 | 0.063246 |  | 0.033912 | 0.020656 | 0.032863 | 0.05164 | 0.015055 | 0.040833 | 0
| Rel.Std.Dev. |  |  |  | 2.877% | 3.218% | 1.21% | 1.829% | 0% | 1.387% | 2.168% | 12.65% |  | 7.621% | 8.154% | 5.666% | 11.07% | 3.795% | 5.697% | 0%
| PDM3 |  |  |  | 17.47% | 19.39% | -2.041% | -5.485% | -5.867% | -9.311% | -13.52% | 14.8% |  | 2.168% | -41.84% | 33.16% | 7.143% | -8.929% | 64.56% | 14.8%
| Z-Score (Absolute) |  | 14.52 | 1.18 | 0.65 | 0.7 | 0.05 | 0.05 | 0.06 | 0.17 | 0.3 | 0.56 | 14.52 | 0.18 | 1.16 | 1.12 | NA | 0.16 | 2.08 | NA | 1.18 | 1.18
| Analytical results for Bi in OREAS 931b (Certified Value 210 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | 0
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 0
| 1 | 1 | 212 | 212.3045 | 212.49 | 187 | 213 | 211 | 221 | 199.5 | 201 | 245.1 | 189.932 | 199 | 211.98 | 234 | 210.37 | 217.8 | 185 | 222.1 | 318.2865 | 190 | 146 | 149 | 200.4788 | 1
| 1 | 2 | 214 | 215.8472 | 214.14 | 185 | 216 | 218 | 211 | 216 | 194.5 | 238.82 | 171.76 | 190 | 242.04 | 232 | 204.49 | 224.4 | 183 | 221.02 | 211.98 | 188 | 146 | 173 | 206.3121 | 8
| 1 | 3 | 213 | 215.476 | 215.63 | 186 | 215 | 217 | 222 | 201 | 199 | 245.2 | 181.961 | 206 | 262.02 | 232 | 211.29 | 207.4 | 188 | 221.25 | 334.0871 | 193 | 142 | 177 | 217.5668 | 16
| 1 | 4 | 220 | 216.6299 | 218.74 | 188 | 212 | 208 | 223 | 207 | 202 | 233.53 | 199.211 | 210 | 223.03 | 236 | 205.9 | 209.1 | 182 | 221.86 | 316.9498 | 196 | 127 | 161 | 209.0035 | 209.9302
| 1 | 5 | 214 | 215.9471 | 214.44 | 184 | 214 | 211 | 219 | 213 | 204 | 238.71 | 191.195 | 194 | 239.13 | 232 | 198.53 | 212.2 | 194 | 218.83 | 332.2345 | 184 | 131 | 165 | 209.9726 | 73
| 1 | 6 | 213 | 212.6466 | 217.22 | 189 | 212 | 211 | 239 | 198 | 198 | 242.21 |  | 209 | 238.23 | 229 | 201.67 | 209.9 | 185 | 216.81 | 340.5969 | 192 | 124 | 153 | 214.7231
| Mean |  | 214.3333 | 214.8085 | 215.4433 | 186.5 | 213.6667 | 212.6667 | 222.5 | 205.75 | 199.75 | 240.595 | 186.8118 | 201.3333 | 236.0717 | 232.5 | 205.375 | 213.4667 | 186.1667 | 220.3117 | 309.0225 | 190.5 | 136 | 163 | 209.6762
| Median |  | 213.5 | 215.6616 | 215.035 | 186.5 | 213.5 | 211 | 221.5 | 204 | 200 | 240.515 | 189.932 | 202.5 | 238.68 | 232 | 205.195 | 211.05 | 185 | 221.135 | 325.2605 | 191 | 136.5 | 163 | 209.488
| Std Dev. |  | 2.8752 | 1.8483 | 2.2577 | 1.8708 | 1.633 | 3.9328 | 9.1597 | 7.495 | 3.3429 | 4.4884 | 10.4016 | 8.2865 | 17.1665 | 2.3452 | 4.9311 | 6.4578 | 4.3551 | 2.0705 | 48.4353 | 4.1833 | 9.859 | 10.9545 | 6.0693
| Rel.Std.Dev. |  | 1.341% | 0.8605% | 1.048% | 1.003% | 0.7643% | 1.849% | 4.117% | 3.643% | 1.674% | 1.866% | 5.568% | 4.116% | 7.272% | 1.009% | 2.401% | 3.025% | 2.339% | 0.9398% | 15.67% | 2.196% | 7.249% | 6.721% | 2.895%
| PDM3 |  | 2.097% | 2.324% | 2.626% | -11.16% | 1.78% | 1.304% | 5.988% | -1.991% | -4.849% | 14.61% | -11.01% | -4.095% | 12.45% | 10.75% | -2.17% | 1.685% | -11.32% | 4.945% | 47.2% | -9.256% | -35.22% | -22.36% | -0.121%
| Z-Score (Absolute) |  | 0.1 | 0.13 | 0.17 | 1.56 | 0.06 | 0 | 0.59 | 0.41 | 0.77 | 1.66 | 1.54 | 0.67 | 1.39 | 1.18 | 0.43 | 0.05 | 1.58 | 0.45 | 5.73 | 1.32 | 4.56 | 2.96 | 0.18
| Analytical results for Ca in OREAS 931b (Certified Value 0.079 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 0.07 | 0.078605 | 0.077 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.07 | 0.09 | 0.074521 | 0.081 | 0.08 | 0.07 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1052 | 0.08 | 0.071 | 0.081 | 0.093834 | 1
| 1 | 2 | 0.07 | 0.077783 | 0.078 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.072749 | 0.082 | 0.09 | 0.07 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1052 | 0.08 | 0.079 | 0.081 | 0.094736 | 0
| 1 | 3 | 0.07 | 0.076255 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.075089 | 0.081 | 0.1 | 0.07 | 0.08 | 0.07 | 0.07 | 0.08 | 0.105 | 0.08 | 0.071 | 0.08 | 0.10054 | 16
| 1 | 4 | 0.07 | 0.076078 | 0.079 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.076035 | 0.083 | 0.08 | 0.07 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1027 | 0.08 | 0.071 | 0.08 | 0.096735 | 0.078702
| 1 | 5 | 0.07 | 0.076006 | 0.078 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.07418 | 0.078 | 0.09 | 0.07 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1049 | 0.08 | 0.079 | 0.081 | 0.097434 | 74
| 1 | 6 | 0.07 | 0.074379 | 0.081 | 0.08 | 0.08 | 0.08 | 0.08 | 0.07 | 0.07 | 0.09 | 0.074596 | 0.085 | 0.09 | 0.07 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1054 | 0.08 | 0.071 | 0.08 | 0.095236
| Mean |  | 0.07 | 0.076518 | 0.078833 | 0.08 | 0.08 | 0.08 | 0.08 | 0.078333 | 0.076667 | 0.088333 | 0.074528 | 0.081667 | 0.088333 | 0.07 | 0.08 | 0.078333 | 0.07 | 0.08 | 0.10473 | 0.08 | 0.073667 | 0.0805 | 0.096418
| Median |  | 0.07 | 0.076166 | 0.0785 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.08 | 0.09 | 0.074558 | 0.0815 | 0.09 | 0.07 | 0.08 | 0.08 | 0.07 | 0.08 | 0.1051 | 0.08 | 0.071 | 0.0805 | 0.095985
| Std Dev. |  | 0 | 0.0014872 | 0.001472 | 0 | 0 | 0 | 0 | 0.0040825 | 0.005164 | 0.0040825 | 0.0010842 | 0.0023381 | 0.0075277 | 0 | 0 | 0.0040825 | 0 | 0 | 0.0010113 | 0 | 0.0041312 | 0.00054772 | 0.002408
| Rel.Std.Dev. |  | 0% | 1.944% | 1.867% | 0% | 0% | 0% | 0% | 5.212% | 6.736% | 4.622% | 1.455% | 2.863% | 8.522% | 0% | 0% | 5.212% | 0% | 0% | 0.9656% | 0% | 5.608% | 0.6804% | 2.497%
| PDM3 |  | -11.06% | -2.775% | 0.1669% | 1.649% | 1.649% | 1.649% | 1.649% | -0.4684% | -2.586% | 12.24% | -5.303% | 3.767% | 12.24% | -11.06% | 1.649% | -0.4684% | -11.06% | 1.649% | 33.08% | 1.649% | -6.398% | 2.285% | 22.51%
| Z-Score (Absolute) |  | 4.05 | 1.41 | 0.47 | 0 | 0 | 0 | 0 | 0.67 | 1.35 | 3.37 | 2.21 | 0.67 | 3.37 | 4.05 | 0 | 0.67 | 4.05 | 0 | 10.01 | 0 | 2.49 | 0.2 | 6.64
| Analytical results for Cd in OREAS 931b (Certified Value 0.62 ppm) | [24 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 2
| 1 | 1 | < 0.5 | 0.59336 | 0.629 | 0.55 | 0.45 | 0.54 | 0.56 | 0.55 | 0.52 | 0.84 | < 5 | 0.74 | 0.56 | 0.65 | 0.68 | 0.6 | 0.57 | 0.62 | 0.60724 | 0.72 | < 1 | < 1 |  | 1
| 1 | 2 | < 0.5 | 0.61876 | 0.624 | 0.55 | 0.45 | 0.56 | 0.53 | 0.56 | 0.58 | 0.79 | < 5 | 0.68 | 0.64 | 0.64 | 0.72 | 0.6 | 0.57 | 0.65 | 0.58438 | 0.68 | < 1 | 1 |  | 20
| 1 | 3 | < 0.5 | 0.65295 | 0.646 | 0.52 | 0.42 | 0.54 | 0.54 | 0.58 | 0.58 | 0.79 | < 5 | 0.69 | 0.66 | 0.61 | 0.68 | 0.6 | 0.54 | 0.66 | 0.62223 | 0.75 | < 1 | 1 |  | 16
| 1 | 4 | < 0.5 | 0.64422 | 0.662 | 0.56 | 0.44 | 0.55 | 0.51 | 0.59 | 0.56 | 0.8 | < 5 | 0.71 | 0.6 | 0.63 | 0.71 | 0.6 | 0.49 | 0.66 | 0.61402 | 0.77 | < 1 | 1 |  | 0.61705
| 1 | 5 | < 0.5 | 0.63644 | 0.654 | 0.56 | 0.47 | 0.5 | 0.56 | 0.59 | 0.55 | 0.77 | < 5 | 0.64 | 0.64 | 0.63 | 0.7 | 0.6 | 0.58 | 0.63 | 0.64762 | 0.72 | < 1 | 1 |  | 75
| 1 | 6 | < 0.5 | 0.59091 | 0.658 | 0.56 | 0.43 | 0.55 | 0.49 | 0.58 | 0.61 | 0.84 | < 5 | 0.86 | 0.65 | 0.62 | 0.69 | 0.6 | 0.53 | 0.66 | 0.61018 | 0.75 | < 1 | 1
| Mean |  |  | 0.62277 | 0.6455 | 0.55 | 0.44333 | 0.54 | 0.53167 | 0.575 | 0.56667 | 0.805 |  | 0.72 | 0.625 | 0.63 | 0.69667 | 0.6 | 0.54667 | 0.64667 | 0.61428 | 0.73167 |  | 1
| Median |  |  | 0.6276 | 0.65 | 0.555 | 0.445 | 0.545 | 0.535 | 0.58 | 0.57 | 0.795 |  | 0.7 | 0.64 | 0.63 | 0.695 | 0.6 | 0.555 | 0.655 | 0.6121 | 0.735 |  | 1
| Std Dev. |  |  | 0.026281 | 0.015719 | 0.015492 | 0.017512 | 0.020976 | 0.027869 | 0.016432 | 0.030768 | 0.02881 |  | 0.076158 | 0.037815 | 0.014142 | 0.01633 | 0 | 0.033862 | 0.017512 | 0.020663 | 0.031885 |  | 0
| Rel.Std.Dev. |  |  | 4.22% | 2.435% | 2.817% | 3.95% | 3.884% | 5.242% | 2.858% | 5.43% | 3.579% |  | 10.58% | 6.05% | 2.245% | 2.344% | 0% | 6.194% | 2.708% | 3.364% | 4.358% |  | 0%
| PDM3 |  |  | 0.9275% | 4.61% | -10.87% | -28.15% | -12.49% | -13.84% | -6.815% | -8.165% | 30.46% |  | 16.68% | 1.288% | 2.098% | 12.9% | -2.763% | -11.41% | 4.799% | -0.4498% | 18.57% |  | 62.06%
| Z-Score (Absolute) |  | 3.3 | 0 | 0.2 | 0.64 | 1.59 | 0.73 | 0.81 | 0.42 | 0.5 | 1.61 | 16.63 | 0.86 | 0.02 | 0.06 | 0.65 | NA | 0.67 | 0.21 | 0.08 | 0.96 | 1.09 | 2.6
| Analytical results for Ce in OREAS 931b (Certified Value 49.9 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | [7 empty] | 2
| 1 | 1 | 60.0315 | 47.777 | 56.5 | 53.2 | 46.8 | 49.9 | 46.6 | 50.4 | 56.32 | 50.896 | 50.41 | 42.8 | 64 | 48.3 | 48.16 | 48.98 | [7 empty] | 1
| 1 | 2 | 59.6495 | 48.155 | 56.3 | 50.5 | 46.7 | 50 | 46.6 | 49.9 | 52.42 | 48.998 | 49.76 | 42.2 | 67 | 46.6 | 49.46 | 48.69 | [7 empty] | 9
| 1 | 3 | 60.1552 | 48.255 | 55.5 | 50.4 | 48 | 47.9 | 47.9 | 50.3 | 50.44 | 51.062 | 53.02 | 42.4 | 64 | 47.8 | 50.42 | 48.05 | [7 empty] | 16
| 1 | 4 | 59.6227 | 48.292 | 57 | 51.2 | 47.6 | 48.1 | 47.1 | 51.6 | 50.98 | 51.025 | 52.37 | 45.3 | 65 | 46.3 | 49.69 | 51.25 | [7 empty] | 49.9306
| 1 | 5 | 60.4248 | 47.989 | 56.6 | 50.9 | 48.8 | 49.6 | 46.6 | 49.2 | 51.29 | 49.621 | 50.24 | 44 | 65 | 48.3 | 48.03 | 48.42 | [7 empty] | 76
| 1 | 6 | 59.4095 | 47.539 | 56.4 | 52.6 | 47.7 | 54.4 | 46.8 | 50 | 52.69 | 50.487 | 53.66 | 42.2 | 65 | 46.3 | 50.29 | 50.31
| Mean |  | 59.8822 | 48.0012 | 56.3833 | 51.4667 | 47.6 | 49.9833 | 46.9333 | 50.2333 | 52.3567 | 50.3482 | 51.5767 | 43.15 | 65 | 47.2667 | 49.3417 | 49.2833
| Median |  | 59.8405 | 48.072 | 56.45 | 51.05 | 47.65 | 49.75 | 46.7 | 50.15 | 51.855 | 50.6915 | 51.39 | 42.6 | 65 | 47.2 | 49.575 | 48.835
| Std Dev. |  | 0.3837 | 0.29553 | 0.49565 | 1.1622 | 0.7823 | 2.3473 | 0.51251 | 0.79162 | 2.1227 | 0.85309 | 1.6432 | 1.2518 | 1.0954 | 0.97297 | 1.0309 | 1.2355
| Rel.Std.Dev. |  | 0.6408% | 0.6157% | 0.8791% | 2.258% | 1.643% | 4.696% | 1.092% | 1.576% | 4.054% | 1.694% | 3.186% | 2.901% | 1.685% | 2.058% | 2.089% | 2.507%
| PDM3 |  | 19.93% | -3.864% | 12.92% | 3.076% | -4.668% | 0.1055% | -6.003% | 0.6062% | 4.859% | 0.8362% | 3.297% | -13.58% | 30.18% | -5.335% | -1.18% | -1.296%
| Z-Score (Absolute) |  | 3.03 | 0.65 | 1.94 | 0.42 | 0.78 | 0.04 | 0.98 | 0.04 | 0.7 | 0.07 | 0.45 | 2.15 | 4.61 | 0.88 | 0.24 | 0.26
| Analytical results for Co in OREAS 931b (Certified Value 29.6 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 2
| 1 | 1 | 32 | 28.0771 | 29.83 | 29.7 | 30.8 | 28.9 | 31.3 | 26.9 | 26.6 | 34.5 | 28.305 | 27.9 | 29.2 | 30.5 | 29 | 29.1 | 27.4 | 29.9 | 26.1353 | 33.1 | 28 | 28 |  | 1
| 1 | 2 | 32 | 28.8868 | 30.13 | 29.2 | 30.2 | 28.7 | 31.5 | 27.3 | 28.8 | 33.8 | 26.855 | 27 | 32.9 | 30.2 | 30 | 29 | 26.4 | 31.2 | 25.967 | 31.1 | 29 | 28 |  | 11
| 1 | 3 | 32 | 28.7998 | 30.34 | 29 | 30.3 | 29.3 | 31.2 | 27.4 | 27.7 | 34.4 | 28.459 | 28.7 | 34.7 | 30.6 | 29 | 28.5 | 27.3 | 31.4 | 26.6766 | 33.5 | 27 | 29 |  | 16
| 1 | 4 | 32 | 28.2046 | 29.85 | 29.4 | 31.3 | 29 | 31.1 | 28.5 | 27.8 | 33.6 | 28.532 | 29 | 30 | 30.9 | 29 | 28.1 | 25.4 | 31 | 26.0364 | 34 | 27 | 29 |  | 29.5683
| 1 | 5 | 32 | 28.8151 | 29.98 | 29.5 | 30.3 | 29.8 | 31.1 | 27.2 | 28.2 | 33.5 | 28.199 | 27.3 | 31.9 | 31 | 29 | 28.8 | 26.7 | 29.9 | 26.6602 | 32.9 | 28 | 29 |  | 77
| 1 | 6 | 32 | 28.6516 | 29.81 | 29.5 | 29.8 | 29.5 | 27.8 | 28.7 | 26.6 | 33.5 | 28.591 | 28.9 | 32.4 | 30 | 30 | 28.9 | 26.3 | 30.4 | 26.586 | 33.7 | 28 | 29
| Mean |  | 32 | 28.5725 | 29.99 | 29.3833 | 30.45 | 29.2 | 30.6667 | 27.6667 | 27.6167 | 33.8833 | 28.1568 | 28.1333 | 31.85 | 30.5333 | 29.3333 | 28.7333 | 26.5833 | 30.6333 | 26.3436 | 33.05 | 27.8333 | 28.6667
| Median |  | 32 | 28.7257 | 29.915 | 29.45 | 30.3 | 29.15 | 31.15 | 27.35 | 27.75 | 33.7 | 28.382 | 28.3 | 32.15 | 30.55 | 29 | 28.85 | 26.55 | 30.7 | 26.3607 | 33.3 | 28 | 29
| Std Dev. |  | 0 | 0.34534 | 0.20967 | 0.24833 | 0.5244 | 0.40988 | 1.4123 | 0.74476 | 0.87731 | 0.4535 | 0.65407 | 0.85946 | 1.9987 | 0.38816 | 0.5164 | 0.37238 | 0.73598 | 0.65929 | 0.33151 | 1.0349 | 0.75277 | 0.5164
| Rel.Std.Dev. |  | 0% | 1.209% | 0.6991% | 0.8451% | 1.722% | 1.404% | 4.605% | 2.692% | 3.177% | 1.338% | 2.323% | 3.055% | 6.276% | 1.271% | 1.76% | 1.296% | 2.769% | 2.152% | 1.258% | 3.131% | 2.705% | 1.801%
| PDM3 |  | 8.224% | -3.368% | 1.426% | -0.6257% | 2.982% | -1.246% | 3.715% | -6.431% | -6.601% | 14.59% | -4.774% | -4.853% | 7.717% | 3.264% | -0.7948% | -2.824% | -10.1% | 3.602% | -10.91% | 11.77% | -5.868% | -3.049%
| Z-Score (Absolute) |  | 1.4 | 0.36 | 0.37 | 0.06 | 0.61 | 0.03 | 0.72 | 0.82 | 0.85 | 2.36 | 0.57 | 0.58 | 1.32 | 0.65 | 0.03 | 0.27 | 1.37 | 0.7 | 1.5 | 1.94 | 0.73 | 0.31
| Analytical results for Cr in OREAS 931b (Certified Value 21.1 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 2
| 1 | 1 | 24 | 20.969 | 21.6 | 24 | 21 | 21 | 21 | 19 | 19 | 23 | 21.293 | 19 | 19.2 | 21 | 25 | 21 | 22 | 21 | 22.805 | 21 | 23 | 20 |  | 1
| 1 | 2 | 23 | 21.26 | 21.3 | 25 | 21 | 21 | 21 | 19 | 20 | 21 | 19.609 | 20 | 22.1 | 20.7 | 26 | 22 | 21 | 20 | 24.2099 | 20 | 22 | 20 |  | 0
| 1 | 3 | 23 | 21.4272 | 21.7 | 25 | 21 | 21 | 21 | 20 | 20 | 22 | 21.269 | 19 | 23 | 20.8 | 24 | 21 | 21 | 20 | 26.4322 | 20 | 21 | 21 |  | 16
| 1 | 4 | 23 | 21.68 | 21.4 | 25 | 21 | 21 | 21 | 20 | 19 | 21 | 21.092 | 20 | 20.1 | 21.4 | 25 | 20 | 20 | 21 | 24.2391 | 21 | 21 | 21 |  | 21.0864
| 1 | 5 | 24 | 21.8994 | 21.3 | 25 | 21 | 21 | 22 | 19 | 19 | 22 | 21.611 | 19 | 21.4 | 21.1 | 25 | 22 | 21 | 20 | 24.6077 | 20 | 21 | 20 |  | 78
| 1 | 6 | 23 | 21.3402 | 21.4 | 26 | 21 | 21 | 21 | 19 | 19 | 21 | 21.066 | 20 | 22.1 | 21 | 25 | 21 | 21 | 20 | 24.0346 | 20 | 21 | 22
| Mean |  | 23.3333 | 21.4293 | 21.45 | 25 | 21 | 21 | 21.1667 | 19.3333 | 19.3333 | 21.6667 | 20.99 | 19.5 | 21.3167 | 21 | 25 | 21.1667 | 21 | 20.3333 | 24.3881 | 20.3333 | 21.5 | 20.6667
| Median |  | 23 | 21.3837 | 21.4 | 25 | 21 | 21 | 21 | 19 | 19 | 21.5 | 21.1805 | 19.5 | 21.75 | 21 | 25 | 21 | 21 | 20 | 24.2245 | 20 | 21 | 20.5
| Std Dev. |  | 0.5164 | 0.3263 | 0.16432 | 0.63246 | 0 | 0 | 0.40825 | 0.5164 | 0.5164 | 0.8165 | 0.70408 | 0.54772 | 1.4162 | 0.24495 | 0.63246 | 0.75277 | 0.63246 | 0.5164 | 1.1757 | 0.5164 | 0.83666 | 0.8165
| Rel.Std.Dev. |  | 2.213% | 1.523% | 0.766% | 2.53% | 0% | 0% | 1.929% | 2.671% | 2.671% | 3.768% | 3.354% | 2.809% | 6.644% | 1.166% | 2.53% | 3.556% | 3.012% | 2.54% | 4.821% | 2.54% | 3.891% | 3.951%
| PDM3 |  | 10.66% | 1.626% | 1.724% | 18.56% | -0.4096% | -0.4096% | 0.3808% | -8.314% | -8.314% | 2.752% | -0.457% | -7.523% | 1.092% | -0.4096% | 18.56% | 0.3808% | -0.4096% | -3.571% | 15.66% | -3.571% | 1.962% | -1.99%
| Z-Score (Absolute) |  | 3.64 | 0.56 | 0.59 | 6.34 | 0.13 | 0.13 | 0.13 | 2.83 | 2.83 | 0.94 | 0.15 | 2.56 | 0.38 | 0.13 | 6.34 | 0.13 | 0.13 | 1.21 | 5.35 | 1.21 | 0.67 | 0.67
| Analytical results for Cs in OREAS 931b (Certified Value 2.06 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | [7 empty] | 3
| 1 | 1 | 2.054 | 2.08 | 2.26 | 1.93 | 1.93 | 1.98 | 1.82 | 1.76 | 2.27 | 2 | 2.2 | 2.54 | 1.92 | 2.16 | 2.1951 | 1.94 | [7 empty] | 1
| 1 | 2 | 2.0537 | 2.07 | 2.21 | 1.81 | 1.94 | 2.03 | 1.85 | 1.81 | 2.22 | 1.93 | 2.21 | 2.5 | 1.85 | 2.23 | 1.806 | 1.89 | [7 empty] | 0
| 1 | 3 | 2.0752 | 2.06 | 2.33 | 1.86 | 1.99 | 1.94 | 1.81 | 1.83 | 2.25 | 2.05 | 2.2 | 2.61 | 1.92 | 2.25 | 2.194 | 1.97 | [7 empty] | 16
| 1 | 4 | 2.0751 | 2.05 | 2.34 | 1.88 | 1.98 | 1.95 | 1.84 | 1.81 | 2.32 | 2.13 | 2.21 | 2.45 | 1.78 | 2.21 | 2.2365 | 2.07 | [7 empty] | 2.0636
| 1 | 5 | 2.0552 | 2.07 | 2.39 | 1.86 | 1.97 | 1.99 | 1.84 | 1.81 | 2.33 | 1.99 | 2.22 | 2.28 | 1.89 | 2.16 | 1.9604 | 2.03 | [7 empty] | 79
| 1 | 6 | 2.0441 | 2.07 | 2.31 | 1.9 | 1.97 | 1.92 | 1.74 | 1.8 | 2.3 | 2.19 | 2.08 | 2.43 | 1.83 | 2.2 | 2.103 | 2.03
| Mean |  | 2.0595 | 2.0667 | 2.3067 | 1.8733 | 1.9633 | 1.9683 | 1.8167 | 1.8033 | 2.2817 | 2.0483 | 2.1867 | 2.4683 | 1.865 | 2.2017 | 2.0825 | 1.9883
| Median |  | 2.0546 | 2.07 | 2.32 | 1.87 | 1.97 | 1.965 | 1.83 | 1.81 | 2.285 | 2.025 | 2.205 | 2.475 | 1.87 | 2.205 | 2.1485 | 2
| Std Dev. |  | 0.012721 | 0.010328 | 0.063456 | 0.040825 | 0.023381 | 0.039707 | 0.040332 | 0.023381 | 0.042622 | 0.096419 | 0.052789 | 0.11268 | 0.055408 | 0.03656 | 0.16768 | 0.067057
| Rel.Std.Dev. |  | 0.6177% | 0.4997% | 2.751% | 2.179% | 1.191% | 2.017% | 2.22% | 1.297% | 1.868% | 4.707% | 2.414% | 4.565% | 2.971% | 1.661% | 8.052% | 3.373%
| PDM3 |  | -0.1952% | 0.1503% | 11.78% | -9.219% | -4.857% | -4.615% | -11.96% | -12.61% | 10.57% | -0.7381% | 5.966% | 19.62% | -9.622% | 6.692% | 0.9174% | -3.646%
| Z-Score (Absolute) |  | 0.03 | 0.06 | 1.22 | 0.87 | 0.44 | 0.41 | 1.14 | 1.21 | 1.1 | 0.03 | 0.64 | 1.99 | 0.91 | 0.71 | 0.14 | 0.32
| Analytical results for Cu in OREAS 931b (Certified Value 3.95 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | AR*AAS | AR*OES | AR*MS | AR*OES/MS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | 2
|  |  | 0.2g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.4g |  | 0.25g | 01g | 0.2g | 0.5g | 3
| 1 | 1 | 4.14 | 3.9125 | > 2 | 3.79 | 3.91 | 3.85 | 3.84 | 3.82 | 3.79 | 4.0856 | 4.1409 | 4.036 | 3.4926 | 3.95 | > 1 | 3.827 | 4.02 | 4.103 | 3.408 | 3.92 | 4.1 | 3.83 | 3.9178 | 1
| 1 | 2 | 4.13 | 4.0008 | > 2 | 3.83 | 3.9 | 3.84 | 3.8 | 3.83 | 3.8 | 4.0336 | 3.9818 | 4.094 | 3.8835 | 3.93 | > 1 | 3.856 | 3.92 | 4.037 | 3.512 | 3.91 | 4.08 | 3.83 | 3.9599 | 0
| 1 | 3 | 4.12 | 3.9407 | > 2 | 3.78 | 3.88 | 3.86 | 3.86 | 3.9 | 3.82 | 4.1205 | 4.0047 | 4.075 | 4.0903 | 3.98 | > 1 | 3.801 | 3.95 | 3.993 | 3.561 | 3.92 | 4.11 | 3.86 | 4.0286 | 16
| 1 | 4 | 4.13 | 3.9145 | > 2 | 3.75 | 3.89 | 3.87 | 3.77 | 3.93 | 3.75 | 3.9953 | 4.2105 | 4.1643 | 3.5951 | 4.03 | > 1 | 3.834 | 4.02 | 3.924 | 3.494 | 3.93 | 4.08 | 3.92 | 4.0363 | 3.9481
| 1 | 5 | 4.14 | 4.0195 | > 2 | 3.85 | 3.9 | 3.87 | 3.92 | 3.89 | 3.84 | 4.084 | 4.1705 | 3.9323 | 3.7884 | 4.03 | > 1 | 3.811 | 4.01 | 3.983 | 3.433 | 4.03 | 4.12 | 3.91 | 4.0602 | 80
| 1 | 6 | 4.13 | 3.9595 | > 2 | 3.81 | 3.91 | 3.87 | 3.67 | 3.89 | 3.82 | 4.0896 | 4.0132 | 4.2208 | 3.8496 | 3.93 | > 1 | 3.855 | 3.98 | 4.036 | 3.471 | 4.02 | 4.11 | 3.99 | 3.9473
| Mean |  | 4.1317 | 3.9579 |  | 3.8017 | 3.8983 | 3.86 | 3.81 | 3.8767 | 3.8033 | 4.0681 | 4.0869 | 4.0871 | 3.7833 | 3.975 |  | 3.8307 | 3.9833 | 4.0127 | 3.4798 | 3.955 | 4.1 | 3.89 | 3.9917
| Median |  | 4.13 | 3.9501 |  | 3.8 | 3.9 | 3.865 | 3.82 | 3.89 | 3.81 | 4.0848 | 4.077 | 4.0845 | 3.819 | 3.965 |  | 3.8305 | 3.995 | 4.0145 | 3.4825 | 3.925 | 4.105 | 3.885 | 3.9942
| Std Dev. |  | 0.0075277 | 0.044466 |  | 0.036009 | 0.01169 | 0.012649 | 0.08579 | 0.04274 | 0.031411 | 0.045286 | 0.098412 | 0.10059 | 0.21382 | 0.046368 |  | 0.022474 | 0.041312 | 0.060705 | 0.055236 | 0.054681 | 0.016733 | 0.06229 | 0.057414
| Rel.Std.Dev. |  | 0.1822% | 1.123% |  | 0.9472% | 0.2999% | 0.3277% | 2.252% | 1.102% | 0.8259% | 1.113% | 2.408% | 2.461% | 5.652% | 1.166% |  | 0.5867% | 1.037% | 1.513% | 1.587% | 1.383% | 0.4081% | 1.601% | 1.438%
| PDM3 |  | 4.65% | 0.2491% |  | -3.708% | -1.26% | -2.231% | -3.497% | -1.809% | -3.666% | 3.04% | 3.517% | 3.521% | -4.174% | 0.6821% |  | -2.974% | 0.8932% | 1.636% | -11.86% | 0.1756% | 3.848% | -1.471% | 1.105%
| Z-Score (Absolute) |  | 1.05 | 0.02 | NA | 0.91 | 0.34 | 0.57 | 0.86 | 0.47 | 0.9 | 0.67 | 0.79 | 0.79 | 1.02 | 0.12 | NA | 0.74 | 0.17 | 0.34 | 2.83 | 0 | 0.86 | 0.39 | 0.22
| Analytical results for Dy in OREAS 931b (Indicative Value 1.68 ppm) | [24 empty] | Indicative
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 15; 03: 17 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g | 2
| 1 | 1 | 01: 1.362; 02: 2.13; 03: 1.6 | 1
| 1 | 2 | 01: 1.401; 02: 2.08; 03: 1.6 | 4
| 1 | 3 | 01: 1.34; 02: 2.13; 03: 1.6 | 16
| 1 | 4 | 01: 1.379; 02: 2.05; 03: 1.6 | 1.6771
| 1 | 5 | 01: 1.371; 02: 1.94; 03: 1.6 | 10
| 1 | 6 | 01: 1.385; 02: 2.02; 03: 1.6
| Mean |  | 01: 1.373; 02: 2.0583; 03: 1.6
| Median |  | 01: 1.375; 02: 2.065; 03: 1.6
| Std Dev. |  | 01: 0.020871; 02: 0.072503; 03: 2.4324e-16
| Rel.Std.Dev. |  | 01: 1.52%; 02: 3.522%; 03: 1.52e-14%
| PDM3 |  | 01: -18.13%; 02: 22.73%; 03: -4.598%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.36; 03: 0
| Analytical results for Er in OREAS 931b (Indicative Value 0.67 ppm) | [24 empty] | Indicative
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 15; 03: 17 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g | 2
| 1 | 1 | 01: 0.554; 02: 0.81; 03: 0.6 | 1
| 1 | 2 | 01: 0.587; 02: 0.82; 03: 0.6 | 5
| 1 | 3 | 01: 0.567; 02: 0.82; 03: 0.6 | 16
| 1 | 4 | 01: 0.597; 02: 0.82; 03: 0.6 | 0.66517
| 1 | 5 | 01: 0.574; 02: 0.75; 03: 0.7 | 11
| 1 | 6 | 01: 0.574; 02: 0.8; 03: 0.6
| Mean |  | 01: 0.5755; 02: 0.80333; 03: 0.61667
| Median |  | 01: 0.574; 02: 0.815; 03: 0.6
| Std Dev. |  | 01: 0.015057; 02: 0.027325; 03: 0.040825
| Rel.Std.Dev. |  | 01: 2.616%; 02: 3.401%; 03: 6.62%
| PDM3 |  | 01: -13.48%; 02: 20.77%; 03: -7.291%
| Z-Score (Absolute) |  | 01: 0.67; 02: 3.06; 03: 0
| Analytical results for Eu in OREAS 931b (Indicative Value 0.6 ppm) | [24 empty] | Indicative
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 15; 03: 17 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g | 2
| 1 | 1 | 01: 0.593; 02: 0.91; 03: 0.6 | 1
| 1 | 2 | 01: 0.615; 02: 0.91; 03: 0.6 | 6
| 1 | 3 | 01: 0.592; 02: 0.94; 03: 0.6 | 16
| 1 | 4 | 01: 0.603; 02: 0.91; 03: 0.6 | 0.60225
| 1 | 5 | 01: 0.613; 02: 0.87; 03: 0.6 | 12
| 1 | 6 | 01: 0.611; 02: 0.89; 03: 0.6
| Mean |  | 01: 0.6045; 02: 0.905; 03: 0.6
| Median |  | 01: 0.607; 02: 0.91; 03: 0.6
| Std Dev. |  | 01: 0.010154; 02: 0.023452; 03: 0
| Rel.Std.Dev. |  | 01: 1.68%; 02: 2.591%; 03: 0%
| PDM3 |  | 01: 0.3736%; 02: 50.27%; 03: -0.3736%
| Z-Score (Absolute) |  | 01: 0; 02: 45.03; 03: 0.67
| Analytical results for Fe in OREAS 931b (Certified Value 8.29 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 8.69 | 8.5752 | 8.149 | 8.15 | 8.15 | 8.27 | 8.49 | 8.01 | 8.19 | 9.24 | 9.2825 | 8.48 | 7.3 | 7.83 | 7.88 | 7.8 | 8.35 | 8.98 | 8.709 | 8.53 | 8.253 | 8.37 | 8.0794 | 1
| 1 | 2 | 8.69 | 8.6211 | 8.262 | 7.99 | 8.33 | 8.31 | 8.51 | 8.09 | 8.2 | 9.18 | 10.2442 | 8.58 | 8.14 | 7.58 | 8.27 | 8.2 | 8.07 | 8.88 | 8.909 | 8.37 | 8.533 | 8.16 | 8.1091 | 0
| 1 | 3 | 8.67 | 8.4921 | 8.329 | 8.08 | 8.2 | 8.45 | 8.57 | 8.05 | 8.21 | 9.27 | 9.2833 | 8.55 | 8.55 | 7.62 | 8.04 | 7.81 | 8.23 | 8.85 | 8.771 | 8.39 | 7.974 | 8.18 | 8.2728 | 16
| 1 | 4 | 8.69 | 8.532 | 8.273 | 8.19 | 8.28 | 8.35 | 8.64 | 8.13 | 8.2 | 9.06 | 9.719 | 8.7 | 7.48 | 7.85 | 8.02 | 7.9 | 7.77 | 8.78 | 8.73 | 8.38 | 7.904 | 8.4 | 8.2711 | 8.2904
| 1 | 5 | 8.71 | 8.4929 | 8.328 | 8.19 | 8.15 | 8.36 | 8.52 | 8.08 | 8.17 | 9.27 | 9.7862 | 8.21 | 7.98 | 7.85 | 8.07 | 8.15 | 8.1 | 8.83 | 8.776 | 8.39 | 8.463 | 8.17 | 8.3015 | 81
| 1 | 6 | 8.71 | 8.5917 | 8.309 | 8.15 | 8.1 | 8.48 | 8.12 | 7.92 | 7.86 | 9.29 | 10.1904 | 8.87 | 8.05 | 7.73 | 8.28 | 7.96 | 7.93 | 8.74 | 8.72 | 8.3 | 8.044 | 8.08 | 8.0542
| Mean |  | 8.6933 | 8.5508 | 8.275 | 8.125 | 8.2017 | 8.37 | 8.475 | 8.0467 | 8.1383 | 9.2183 | 9.7509 | 8.565 | 7.9167 | 7.7433 | 8.0933 | 7.97 | 8.075 | 8.8433 | 8.7692 | 8.3933 | 8.1952 | 8.2267 | 8.1814
| Median |  | 8.69 | 8.5536 | 8.291 | 8.15 | 8.175 | 8.355 | 8.515 | 8.065 | 8.195 | 9.255 | 9.7526 | 8.565 | 8.015 | 7.78 | 8.055 | 7.93 | 8.085 | 8.84 | 8.7505 | 8.385 | 8.1485 | 8.175 | 8.1901
| Std Dev. |  | 0.015055 | 0.053591 | 0.06772 | 0.077395 | 0.087502 | 0.080747 | 0.18207 | 0.073937 | 0.13703 | 0.086583 | 0.41876 | 0.22134 | 0.45689 | 0.12028 | 0.15513 | 0.17018 | 0.20705 | 0.083586 | 0.073722 | 0.075011 | 0.26297 | 0.12801 | 0.11192
| Rel.Std.Dev. |  | 0.1732% | 0.6267% | 0.8184% | 0.9526% | 1.067% | 0.9647% | 2.148% | 0.9189% | 1.684% | 0.9393% | 4.295% | 2.584% | 5.771% | 1.553% | 1.917% | 2.135% | 2.564% | 0.9452% | 0.8407% | 0.8937% | 3.209% | 1.556% | 1.368%
| PDM3 |  | 4.86% | 3.142% | -0.1857% | -1.995% | -1.07% | 0.9602% | 2.227% | -2.94% | -1.834% | 11.19% | 17.62% | 3.312% | -4.508% | -6.599% | -2.377% | -3.865% | -2.598% | 6.67% | 5.775% | 1.242% | -1.149% | -0.7687% | -1.315%
| Z-Score (Absolute) |  | 1.75 | 1.21 | 0.18 | 0.38 | 0.09 | 0.54 | 0.93 | 0.67 | 0.33 | 3.71 | 5.71 | 1.27 | 1.16 | 1.81 | 0.5 | 0.96 | 0.57 | 2.31 | 2.03 | 0.62 | 0.12 | 0 | 0.17
| Analytical results for Ga in OREAS 931b (Certified Value 3.25 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS |  |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g |  |  |  | 2
| 1 | 1 | 8 | 3.3133 | 3 | 6 | 3.16 | 3.03 | 3.13 | 2.93 | 2.71 | 2.6 | 5.496 | 4.5 | 3.4 | 3.38 | 4.02 | 2 | 3.45 | 3.52 | 3.3445 | 3.6 |  |  |  | 1
| 1 | 2 | 8 | 3.4811 | 3 | 5.9 | 3.1 | 3.09 | 3.1 | 2.93 | 2.7 | 2.4 | 5.38 | 4.4 | 3.8 | 3.33 | 4.12 | 3 | 3.36 | 3.62 | 3.104 | 3.5 |  |  |  | 14
| 1 | 3 | 8 | 3.4268 | 3.1 | 6 | 2.98 | 3.16 | 3.06 | 2.85 | 2.78 | 2.5 | 5.299 | 4.7 | 4 | 3.29 | 4.06 | 2 | 3.33 | 3.65 | 3.3077 | 3.7 |  |  |  | 16
| 1 | 4 | 8 | 3.4454 | 3.1 | 5.7 | 3.13 | 3.06 | 3.07 | 2.87 | 2.95 | 2.3 | 5.591 | 4.7 | 3.5 | 3.37 | 4.05 | 2 | 3.28 | 3.62 | 3.3406 | 3.7 |  |  |  | 3.2486
| 1 | 5 | 8 | 3.3233 | 3.1 | 5.9 | 3.05 | 3.13 | 3.09 | 2.95 | 2.84 | 2.4 | 5.394 | 4.5 | 3.6 | 3.36 | 4.01 | 3 | 3.43 | 3.47 | 3.1962 | 3.7 |  |  |  | 82
| 1 | 6 | 8 | 3.3467 | 3.1 | 6 | 3.06 | 3.18 | 3.03 | 2.79 | 2.81 | 2.4 | 5.296 | 4.9 | 3.7 | 3.23 | 4.06 | 3 | 3.4 | 3.57 | 3.1431 | 3.7
| Mean |  | 8 | 3.3894 | 3.0667 | 5.9167 | 3.08 | 3.1083 | 3.08 | 2.8867 | 2.7983 | 2.4333 | 5.4093 | 4.6167 | 3.6667 | 3.3267 | 4.0533 | 2.5 | 3.375 | 3.575 | 3.2393 | 3.65
| Median |  | 8 | 3.3867 | 3.1 | 5.95 | 3.08 | 3.11 | 3.08 | 2.9 | 2.795 | 2.4 | 5.387 | 4.6 | 3.65 | 3.345 | 4.055 | 2.5 | 3.38 | 3.595 | 3.2519 | 3.7
| Std Dev. |  | 0 | 0.070606 | 0.05164 | 0.1169 | 0.064187 | 0.058452 | 0.034641 | 0.06121 | 0.092394 | 0.10328 | 0.11544 | 0.18348 | 0.21602 | 0.057504 | 0.038816 | 0.54772 | 0.064109 | 0.06892 | 0.10527 | 0.083666
| Rel.Std.Dev. |  | 0% | 2.083% | 1.684% | 1.976% | 2.084% | 1.881% | 1.125% | 2.12% | 3.302% | 4.244% | 2.134% | 3.974% | 5.892% | 1.729% | 0.9576% | 21.91% | 1.9% | 1.928% | 3.25% | 2.292%
| PDM3 |  | 146.3% | 4.335% | -5.6% | 82.13% | -5.189% | -4.317% | -5.189% | -11.14% | -13.86% | -25.1% | 66.51% | 42.11% | 12.87% | 2.404% | 24.77% | -23.04% | 3.891% | 10.05% | -0.2844% | 12.36%
| Z-Score (Absolute) |  | NA | 0.09 | 0.66 | 5.93 | 0.63 | 0.56 | 0.63 | 1.07 | 1.28 | 2.12 | 4.76 | 2.93 | 0.73 | 0.06 | 1.62 | NA | 0.06 | 0.52 | 0.26 | 0.69
Results from laboratories 1 and 16 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 931b (Indicative Value 2.9 ppm) | [24 empty] | Indicative
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  |  |  | 1
| No | No | 01: 03; 02: 15; 03: 17 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS |  |  |  | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g |  |  |  | 2
| 1 | 1 | 01: 2.421; 02: 3.62; 03: 2.6 |  |  |  | 1
| 1 | 2 | 01: 2.487; 02: 3.59; 03: 2.8 |  |  |  | 15
| 1 | 3 | 01: 2.46; 02: 3.7; 03: 2.7 |  |  |  | 16
| 1 | 4 | 01: 2.431; 02: 3.52; 03: 2.7 |  |  |  | 2.8962
| 1 | 5 | 01: 2.452; 02: 3.37; 03: 2.7 |  |  |  | 9
| 1 | 6 | 01: 2.461; 02: 3.52; 03: 2.6
| Mean |  | 01: 2.452; 02: 3.5533; 03: 2.6833
| Median |  | 01: 2.456; 02: 3.555; 03: 2.7
| Std Dev. |  | 01: 0.023546; 02: 0.11237; 03: 0.075277
| Rel.Std.Dev. |  | 01: 0.9603%; 02: 3.162%; 03: 2.805%
| PDM3 |  | 01: -15.34%; 02: 22.69%; 03: -7.351%
| Z-Score (Absolute) |  | 01: 0.67; 02: 2.54; 03: 0
| Analytical results for Ge in OREAS 931b (Certified Value 0.18 ppm) | [24 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 17 | 18 | 19 | 20 | [10 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g |  | 0.25g | [10 empty] | 2
| 1 | 1 | < 0.05 | 0.22 | 0.16 | 0.22 | 0.16 | 0.17 | < 0.1 | < 0.1 | 0.18 | 0.2 | 0.12 | 0.31403 | < 0.1 | [10 empty] | 1
| 1 | 2 | < 0.05 | 0.2 | 0.17 | 0.22 | 0.15 | 0.19 | < 0.1 | < 0.1 | 0.15 | 0.2 | 0.13 | 0.28814 | < 0.1 | [10 empty] | 0
| 1 | 3 | < 0.05 | 0.22 | 0.17 | 0.21 | 0.15 | 0.18 | < 0.1 | < 0.1 | 0.18 | 0.2 | 0.14 | 0.33741 | < 0.1 | [10 empty] | 16
| 1 | 4 | < 0.05 | 0.21 | 0.17 | 0.2 | 0.13 | 0.18 | < 0.1 | < 0.1 | 0.17 | 0.2 | 0.13 | 0.30802 | < 0.1 | [10 empty] | 0.175
| 1 | 5 | < 0.05 | 0.2 | 0.17 | 0.22 | 0.15 | 0.19 | < 0.1 | < 0.1 | 0.18 | 0.2 | 0.14 | 0.31684 | < 0.1 | [10 empty] | 83
| 1 | 6 | < 0.05 | 0.22 | 0.18 | 0.21 | 0.14 | 0.18 | < 0.1 | < 0.1 | 0.15 | 0.2 | 0.14 | 0.34364 | < 0.1
| Mean |  | 02: 0.21167; 03: 0.17; 04: 0.21333; 05: 0.14667; 06: 0.18167; 09: 0.16833; 10: 0.2; 11: 0.13333; 12: 0.31801
| Median |  | 02: 0.215; 03: 0.17; 04: 0.215; 05: 0.15; 06: 0.18; 09: 0.175; 10: 0.2; 11: 0.135; 12: 0.31543
| Std Dev. |  | 02: 0.0098319; 03: 0.0063246; 04: 0.008165; 05: 0.010328; 06: 0.0075277; 09: 0.01472; 10: 3.0405e-17; 11: 0.008165; 12: 0.020211
| Rel.Std.Dev. |  | 02: 4.645%; 03: 3.72%; 04: 3.827%; 05: 7.042%; 06: 4.144%; 09: 8.744%; 10: 1.52e-14%; 11: 6.124%; 12: 6.356%
| PDM3 |  | 02: 20.95%; 03: -2.857%; 04: 21.9%; 05: -16.19%; 06: 3.81%; 09: -3.81%; 10: 14.29%; 11: -23.81%; 12: 81.72%
| Z-Score (Absolute) |  | 1.62 | 0.66 | 0.15 | 0.68 | 0.13 | 0.3 | 1.32 | 1.32 | 0.13 | NA | 0.3 | 1.97 | 1.32
Results from laboratory 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 931b (Certified Value 0.34 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [7 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | [7 empty] | 3
| 1 | 1 | 0.32154 | 0.35 | 0.39 | 0.33 | 0.32 | 0.34 | 0.33 | 0.35 | 0.36 | 0.33 | 0.42 | 0.35 | 0.5 | 0.4 | 0.3418 | 0.33 | [7 empty] | 1
| 1 | 2 | 0.32289 | 0.35 | 0.4 | 0.34 | 0.33 | 0.33 | 0.35 | 0.34 | 0.33 | 0.4 | 0.41 | 0.37 | 0.5 | 0.42 | 0.12246 | 0.33 | [7 empty] | 16
| 1 | 3 | 0.31506 | 0.37 | 0.38 | 0.31 | 0.33 | 0.35 | 0.33 | 0.34 | 0.37 | 0.35 | 0.42 | 0.35 | 0.5 | 0.42 | 0.30962 | 0.33 | [7 empty] | 16
| 1 | 4 | 0.30255 | 0.36 | 0.39 | 0.34 | 0.33 | 0.32 | 0.32 | 0.34 | 0.35 | 0.35 | 0.43 | 0.35 | 0.5 | 0.4 | 0.29978 | 0.34 | [7 empty] | 0.33978
| 1 | 5 | 0.29999 | 0.34 | 0.34 | 0.32 | 0.33 | 0.33 | 0.33 | 0.33 | 0.33 | 0.34 | 0.42 | 0.35 | 0.5 | 0.4 | 0.34341 | 0.32 | [7 empty] | 84
| 1 | 6 | 0.29969 | 0.36 | 0.34 | 0.33 | 0.34 | 0.3 | 0.36 | 0.33 | 0.34 | 0.37 | 0.41 | 0.35 | 0.5 | 0.4 | 0.35602 | 0.33
| Mean |  | 0.31029 | 0.355 | 0.37333 | 0.32833 | 0.33 | 0.32833 | 0.33667 | 0.33833 | 0.34667 | 0.35667 | 0.41833 | 0.35333 | 0.5 | 0.40667 | 0.29552 | 0.33
| Median |  | 0.3088 | 0.355 | 0.385 | 0.33 | 0.33 | 0.33 | 0.33 | 0.34 | 0.345 | 0.35 | 0.42 | 0.35 | 0.5 | 0.4 | 0.32571 | 0.33
| Std Dev. |  | 0.01083 | 0.010488 | 0.026583 | 0.01169 | 0.0063246 | 0.017224 | 0.015055 | 0.0075277 | 0.01633 | 0.025033 | 0.0075277 | 0.008165 | 0 | 0.010328 | 0.087479 | 0.0063246
| Rel.Std.Dev. |  | 3.49% | 2.954% | 7.121% | 3.561% | 1.917% | 5.246% | 4.472% | 2.225% | 4.711% | 7.019% | 1.799% | 2.311% | 0% | 2.54% | 29.6% | 1.917%
| PDM3 |  | -8.679% | 4.481% | 9.877% | -3.368% | -2.877% | -3.368% | -0.9149% | -0.4244% | 2.028% | 4.971% | 23.12% | 3.99% | 47.16% | 19.69% | -13.03% | -2.877%
| Z-Score (Absolute) |  | 1.53 | 0.59 | 1.47 | 0.67 | 0.59 | 0.67 | 0.28 | 0.2 | 0.2 | 0.67 | 3.61 | 0.52 | 7.5 | 3.05 | 2.24 | 0.59
| Analytical results for Hg in OREAS 931b (Certified Value 0.11 ppm) | [24 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | < 5 | 0.14631 | 0.1 | 0.1 | 0.1 | 0.11 | 0.11 | 0.1 | 0.23 | 2.49 | 0.105 | 0.37 | 0.07 | 0.1 | 0.07 | 0.11 | 0.20676 | 0.11 | < 1 | < 1 | 0.1587 |  |  | 1
| 1 | 2 | < 5 | 0.14547 | 0.1 | 0.09 | 0.1 | 0.11 | 0.1 | 0.1 | 0.2 | 2.68 | 0.126 | 0.36 | 0.07 | 0.11 | 0.09 | 0.118 | 0.18833 | 0.13 | < 1 | 1 | 0.1634 |  |  | 1
| 1 | 3 | < 5 | 0.14549 | 0.1 | 0.11 | 0.11 | 0.1 | 0.12 | 0.1 | 0.22 | 2.28 | 0.136 | 0.36 | 0.08 | 0.11 | 0.08 | 0.132 | 0.24028 | 0.13 | < 1 | < 1 | 0.1838 |  |  | 16
| 1 | 4 | < 5 | 0.15383 | 0.1 | 0.1 | 0.1 | 0.11 | 0.11 | 0.12 | 0.21 | 2.2 | 0.126 | 0.35 | 0.1 | 0.09 | 0.08 | 0.123 | 0.27779 | 0.12 | < 1 | < 1 | 0.182 |  |  | 0.11402
| 1 | 5 | < 5 | 0.15171 | 0.1 | 0.08 | 0.13 | 0.11 | 0.11 | 0.12 | 0.27 | 1.86 | 0.136 | 0.68 | 0.07 | 0.09 | 0.11 | 0.113 | 0.23895 | 0.13 | < 1 | < 1 | 0.1828 |  |  | 85
| 1 | 6 | < 5 | 0.14092 | 0.11 | 0.09 | 0.12 | 0.12 | 0.09 | 0.12 | 0.25 | 2.03 | 0.126 | 0.46 | 0.08 | 0.11 | 0.1 | 0.122 | 0.2584 | 0.11 | < 1 | < 1 | 0.1854
| Mean |  |  | 0.14729 | 0.10167 | 0.095 | 0.11 | 0.11 | 0.10667 | 0.11 | 0.23 | 2.2567 | 0.12583 | 0.43 | 0.078333 | 0.10167 | 0.088333 | 0.11967 | 0.23508 | 0.12167 |  | 1 | 0.17602
| Median |  |  | 0.1459 | 0.1 | 0.095 | 0.105 | 0.11 | 0.11 | 0.11 | 0.225 | 2.24 | 0.126 | 0.365 | 0.075 | 0.105 | 0.085 | 0.12 | 0.23961 | 0.125 |  | 1 | 0.1824
| Std Dev. |  |  | 0.0046993 | 0.0040825 | 0.010488 | 0.012649 | 0.0063246 | 0.010328 | 0.010954 | 0.026077 | 0.29871 | 0.011321 | 0.129 | 0.01169 | 0.0098319 | 0.01472 | 0.0078655 | 0.032855 | 0.0098319 |  |  | 0.011743
| Rel.Std.Dev. |  |  | 3.191% | 4.016% | 11.04% | 11.5% | 5.75% | 9.682% | 9.959% | 11.34% | 13.24% | 8.997% | 30% | 14.92% | 9.671% | 16.66% | 6.573% | 13.98% | 8.081% |  |  | 6.672%
| PDM3 |  |  | 29.17% | -10.84% | -16.68% | -3.527% | -3.527% | -6.451% | -3.527% | 101.7% | 1879% | 10.36% | 277.1% | -31.3% | -10.84% | -22.53% | 4.951% | 106.2% | 6.705% |  | 777% | 54.37%
| Z-Score (Absolute) |  | 60.14 | 0.65 | 0.51 | 0.67 | 0.3 | 0.3 | 0.38 | 0.3 | 2.74 | 53.99 | 0.11 | 7.8 | 1.1 | 0.51 | 0.84 | 0.05 | 2.87 | 0 | 9.57 | 11.67 | 1.37
| Analytical results for Ho in OREAS 931b (Indicative Value 0.28 ppm) | [24 empty] | Indicative
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab |  |  | 1
| No | No | 01: 03; 02: 15; 03: 17; 05: 19 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS; 05: AR*MS |  |  | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g |  |  | 2
| 1 | 1 | 01: 0.233; 02: 0.32; 03: 0.3; 05: 0.27388 |  |  | 1
| 1 | 2 | 01: 0.236; 02: 0.33; 03: 0.3; 05: 0.28491 |  |  | 4
| 1 | 3 | 01: 0.226; 02: 0.34; 03: 0.3; 05: 0.28729 |  |  | 16
| 1 | 4 | 01: 0.228; 02: 0.32; 03: 0.3; 05: 0.28718 |  |  | 0.28382
| 1 | 5 | 01: 0.231; 02: 0.3; 03: 0.3; 05: 0.27975 |  |  | 10
| 1 | 6 | 01: 0.227; 02: 0.32; 03: 0.3; 05: 0.28774
| Mean |  | 01: 0.23017; 02: 0.32167; 03: 0.3; 05: 0.28346
| Median |  | 01: 0.2295; 02: 0.32; 03: 0.3; 05: 0.28604
| Std Dev. |  | 01: 0.0038687; 02: 0.013292; 03: 0; 05: 0.0055581
| Rel.Std.Dev. |  | 01: 1.681%; 02: 4.132%; 03: 0%; 05: 1.961%
| PDM3 |  | 01: -18.9%; 02: 13.33%; 03: 5.7%; 05: -0.129%
| Z-Score (Absolute) |  | 01: 2.17; 02: 1.06; 03: 0.29; 05: 0.29
| Analytical results for In in OREAS 931b (Certified Value 3.24 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | [6 empty] | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | [6 empty] | 3
| 1 | 1 | < 5 | 3.1973 | 3.18 | 3.37 | 3.4 | 3.14 | 3.41 | 3.25 | 3.18 | 3.73 | 3.08 | 3.35 | 3.29 | 3.13 | 3.135 | 3.0712 | 3.7 | [6 empty] | 1
| 1 | 2 | < 5 | 3.2209 | 3.16 | 3.38 | 3.37 | 3.18 | 3.37 | 3.3 | 3.23 | 3.76 | 3.226 | 3.34 | 3.25 | 3.1 | 3.23 | 3.0496 | 3.51 | [6 empty] | 0
| 1 | 3 | < 5 | 3.2201 | 3.16 | 3.45 | 3.37 | 3.22 | 3.41 | 3.19 | 3.34 | 3.83 | 3.16 | 3.34 | 3.3 | 3.12 | 3.333 | 2.9784 | 3.68 | [6 empty] | 16
| 1 | 4 | < 5 | 3.2349 | 3.16 | 3.33 | 3.46 | 3.17 | 3.38 | 3.25 | 3.29 | 3.64 | 3.133 | 3.37 | 3.27 | 3.04 | 3.387 | 3.0785 | 3.69 | [6 empty] | 3.2398
| 1 | 5 | < 5 | 3.1884 | 3.1 | 3.45 | 3.37 | 3.16 | 3.42 | 3.14 | 3.36 | 3.66 | 2.955 | 3.37 | 3.21 | 3.14 | 3.154 | 3.0697 | 3.66 | [6 empty] | 86
| 1 | 6 | < 5 | 3.1952 | 3.12 | 3.44 | 3.36 | 3.29 | 3.51 | 3.16 | 3.29 | 3.7 | 3.244 | 3.28 | 3.16 | 3.07 | 3.265 | 3.0504 | 3.63
| Mean |  |  | 3.2095 | 3.1467 | 3.4033 | 3.3883 | 3.1933 | 3.4167 | 3.215 | 3.2817 | 3.72 | 3.133 | 3.3417 | 3.2467 | 3.1 | 3.2507 | 3.0496 | 3.645
| Median |  |  | 3.2087 | 3.16 | 3.41 | 3.37 | 3.175 | 3.41 | 3.22 | 3.29 | 3.715 | 3.1465 | 3.345 | 3.26 | 3.11 | 3.2475 | 3.06 | 3.67
| Std Dev. |  |  | 0.018336 | 0.030111 | 0.050465 | 0.037639 | 0.054283 | 0.049666 | 0.061563 | 0.067355 | 0.06957 | 0.10601 | 0.033116 | 0.053166 | 0.038471 | 0.098713 | 0.036804 | 0.07064
| Rel.Std.Dev. |  |  | 0.5713% | 0.9569% | 1.483% | 1.111% | 1.7% | 1.454% | 1.915% | 2.052% | 1.87% | 3.384% | 0.991% | 1.638% | 1.241% | 3.037% | 1.207% | 1.938%
| PDM3 |  |  | -0.935% | -2.874% | 5.049% | 4.586% | -1.433% | 5.46% | -0.7646% | 1.293% | 14.82% | -3.296% | 3.145% | 0.2129% | -4.314% | 0.3363% | -5.869% | 12.51%
| Z-Score (Absolute) |  | 4.43 | 0.22 | 0.59 | 0.93 | 0.84 | 0.32 | 1.01 | 0.19 | 0.21 | 2.81 | 0.67 | 0.56 | 0 | 0.87 | 0.02 | 1.17 | 2.36
| Analytical results for K in OREAS 931b (Certified Value 0.19 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  | 3
| 1 | 1 | 0.23 | 0.23906 | 0.188 | 0.285 | 0.2 | 0.18 | 0.19 | 0.17 | 0.18 | 0.2 | 0.22427 | 0.182 | 0.22 | 0.16 | 0.2671 | 0.18 | 0.21 | 0.19 | 0.2532 | 0.18 | 0.158 | 0.2 |  | 1
| 1 | 2 | 0.22 | 0.2354 | 0.189 | 0.296 | 0.2 | 0.18 | 0.19 | 0.17 | 0.18 | 0.2 | 0.22134 | 0.187 | 0.26 | 0.16 | 0.2791 | 0.18 | 0.2 | 0.19 | 0.24288 | 0.18 | 0.158 | 0.2 |  | 0
| 1 | 3 | 0.23 | 0.23749 | 0.191 | 0.296 | 0.19 | 0.19 | 0.19 | 0.18 | 0.18 | 0.19 | 0.22394 | 0.185 | 0.26 | 0.16 | 0.276 | 0.17 | 0.19 | 0.19 | 0.25872 | 0.18 | 0.149 | 0.19 |  | 16
| 1 | 4 | 0.22 | 0.21585 | 0.189 | 0.276 | 0.19 | 0.18 | 0.19 | 0.18 | 0.18 | 0.2 | 0.22878 | 0.187 | 0.23 | 0.17 | 0.2705 | 0.17 | 0.2 | 0.18 | 0.26964 | 0.19 | 0.149 | 0.2 |  | 0.19021
| 1 | 5 | 0.22 | 0.22487 | 0.187 | 0.29 | 0.19 | 0.18 | 0.19 | 0.18 | 0.18 | 0.2 | 0.22189 | 0.177 | 0.24 | 0.16 | 0.2668 | 0.18 | 0.21 | 0.18 | 0.23124 | 0.18 | 0.149 | 0.2 |  | 87
| 1 | 6 | 0.23 | 0.2174 | 0.189 | 0.287 | 0.19 | 0.19 | 0.18 | 0.17 | 0.17 | 0.2 | 0.22483 | 0.193 | 0.25 | 0.16 | 0.2681 | 0.18 | 0.2 | 0.18 | 0.2388 | 0.18 | 0.149 | 0.19
| Mean |  | 0.225 | 0.22834 | 0.18883 | 0.28833 | 0.19333 | 0.18333 | 0.18833 | 0.175 | 0.17833 | 0.19833 | 0.22418 | 0.18517 | 0.24333 | 0.16167 | 0.27127 | 0.17667 | 0.20167 | 0.185 | 0.24908 | 0.18167 | 0.152 | 0.19667
| Median |  | 0.225 | 0.23013 | 0.189 | 0.2885 | 0.19 | 0.18 | 0.19 | 0.175 | 0.18 | 0.2 | 0.2241 | 0.186 | 0.245 | 0.16 | 0.2693 | 0.18 | 0.2 | 0.185 | 0.24804 | 0.18 | 0.149 | 0.2
| Std Dev. |  | 0.0054772 | 0.010356 | 0.0013292 | 0.0075542 | 0.005164 | 0.005164 | 0.0040825 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0026441 | 0.0053821 | 0.01633 | 0.0040825 | 0.0051321 | 0.005164 | 0.0075277 | 0.0054772 | 0.014107 | 0.0040825 | 0.0046476 | 0.005164
| Rel.Std.Dev. |  | 2.434% | 4.535% | 0.7039% | 2.62% | 2.671% | 2.817% | 2.168% | 3.13% | 2.289% | 2.058% | 1.179% | 2.907% | 6.711% | 2.525% | 1.892% | 2.923% | 3.733% | 2.961% | 5.664% | 2.247% | 3.058% | 2.626%
| PDM3 |  | 18.29% | 20.05% | -0.7218% | 51.59% | 1.644% | -3.613% | -0.9847% | -7.995% | -6.242% | 4.273% | 17.86% | -2.65% | 27.93% | -15% | 42.62% | -7.118% | 6.025% | -2.737% | 30.95% | -4.49% | -20.09% | 3.397%
| Z-Score (Absolute) |  | 1.68 | 1.85 | 0.11 | 4.83 | 0.11 | 0.38 | 0.14 | 0.8 | 0.63 | 0.36 | 1.64 | 0.29 | 2.59 | 1.46 | 3.98 | 0.72 | 0.53 | 0.3 | 2.88 | 0.47 | 1.93 | 0.28
| Analytical results for La in OREAS 931b (Certified Value 19.5 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  |  | 1
| 1 | 1 | 21 | 23.5658 | 18.091 | 25.8 | 20.6 | 17.2 | 19.7 | 18.2 | 19.5 | 20.9 | 22.3 | 18.9 | 21.1 | 14.9 | 27 | 19 | 20.8 | 18.8 | 18 | 8 | 21.2 |  |  | 1
| 1 | 2 | 22 | 22.8054 | 18.329 | 25.5 | 19.9 | 17.3 | 20.3 | 18 | 19.5 | 19.2 | 21.749 | 18.8 | 22.4 | 14.6 | 28 | 17 | 19.8 | 18.9 | 18.5 | 8 | 21.8 |  |  | 17
| 1 | 3 | 21 | 24.2957 | 18.551 | 25.9 | 19.9 | 17.8 | 18.8 | 18.7 | 20.1 | 18 | 22.248 | 19.8 | 21.9 | 14.9 | 27 | 17 | 20.6 | 17.4 | 18 | 7 | 22.6 |  |  | 16
| 1 | 4 | 21 | 24.5128 | 18.151 | 26.6 | 20 | 17.6 | 18.8 | 18.6 | 19.7 | 18.8 | 22.384 | 19.5 | 20.8 | 15.6 | 27 | 16 | 19.4 | 18 | 19.6 | 7 | 22 |  |  | 19.5236
| 1 | 5 | 21 | 23.3569 | 18.009 | 25.5 | 19.9 | 18.1 | 19.7 | 18.4 | 19.1 | 18.7 | 21.808 | 18.9 | 21.5 | 15.2 | 27 | 16 | 20.5 | 18.7 | 18.3 | 8 | 21.6 |  |  | 88
| 1 | 6 | 21 | 22.1753 | 18.569 | 25.2 | 20.9 | 17.6 | 19.6 | 17.6 | 19.1 | 19.2 | 22.108 | 20.3 | 21.9 | 14.6 | 28 | 17 | 19.7 | 17.5 | 19.4 | 8 | 20.4
| Mean |  | 21.1667 | 23.452 | 18.2833 | 25.75 | 20.2 | 17.6 | 19.4833 | 18.25 | 19.5 | 19.1333 | 22.0995 | 19.3667 | 21.6 | 14.9667 | 27.3333 | 17 | 20.1333 | 18.2167 | 18.6333 | 7.6667 | 21.6
| Median |  | 21 | 23.4614 | 18.24 | 25.65 | 19.95 | 17.6 | 19.65 | 18.3 | 19.5 | 19 | 22.178 | 19.2 | 21.7 | 14.9 | 27 | 17 | 20.15 | 18.35 | 18.4 | 8 | 21.7
| Std Dev. |  | 0.40825 | 0.88396 | 0.23879 | 0.48477 | 0.43818 | 0.32863 | 0.58452 | 0.40866 | 0.37947 | 0.97091 | 0.26499 | 0.60553 | 0.58652 | 0.38297 | 0.5164 | 1.0954 | 0.57155 | 0.67355 | 0.70048 | 0.5164 | 0.74833
| Rel.Std.Dev. |  | 1.929% | 3.769% | 1.306% | 1.883% | 2.169% | 1.867% | 3% | 2.239% | 1.946% | 5.074% | 1.199% | 3.127% | 2.715% | 2.559% | 1.889% | 6.444% | 2.839% | 3.697% | 3.759% | 6.736% | 3.464%
| PDM3 |  | 8.416% | 20.12% | -6.353% | 31.89% | 3.464% | -9.853% | -0.2064% | -6.524% | -0.1211% | -1.999% | 13.19% | -0.804% | 10.64% | -23.34% | 40% | -12.93% | 3.123% | -6.694% | -4.56% | -60.73% | 10.64%
| Z-Score (Absolute) |  | NA | 2.12 | 0.65 | 3.35 | 0.38 | 1.01 | 0 | 0.67 | 0 | 0.19 | 1.4 | 0.07 | 1.13 | 2.42 | 4.2 | NA | 0.34 | 0.68 | 0.46 | NA | 1.13
Results from laboratories 1, 16 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 931b (Certified Value 2.03 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 23 | [5 empty] | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [5 empty] | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 0.25g | [5 empty] | 2
| 1 | 1 | 3 | 2.5864 | 1.66 | 2.6 | 2 | 1.9 | 2 | 2 | 1.8 | < 1 | < 5 | < 1 | < 0.1 | 3.1 | 2.4 | 1.7 | 4 | 3 | [5 empty] | 1
| 1 | 2 | 3 | 2.3863 | 1.62 | 2.6 | 2.1 | 1.9 | 2 | 1.9 | 1.8 | < 1 | < 5 | < 1 | < 0.1 | 3 | 2.4 | 1.6 | 4 | 3 | [5 empty] | 0
| 1 | 3 | 3 | 2.5045 | 1.62 | 2.8 | 2 | 2 | 2 | 1.9 | 1.8 | < 1 | < 5 | < 1 | < 0.1 | 3.2 | 2.5 | 1.5 | 4 | 3 | [5 empty] | 16
| 1 | 4 | 3 | 2.2941 | 1.57 | 2.6 | 2 | 2 | 2 | 2 | 1.8 | < 1 | < 5 | < 1 | < 0.1 | 2.9 | 2.4 | 1.5 | 4 | 3 | [5 empty] | 2.031
| 1 | 5 | 3 | 2.3227 | 1.57 | 2.7 | 2 | 1.9 | 1.9 | 2 | 1.9 | < 1 | < 5 | 4 | < 0.1 | 2.7 | 2.6 | 1.5 | 5 | 4 | [5 empty] | 89
| 1 | 6 | 3 | 2.236 | 1.59 | 2.6 | 1.9 | 2 | 1.8 | 2 | 1.8 | < 1 | < 5 | < 1 | < 0.1 | 2.9 | 2.4 | 1.4 | 5 | 3
| Mean |  | 3 | 2.3883 | 1.605 | 2.65 | 2 | 1.95 | 1.95 | 1.9667 | 1.8167 |  |  | 4 |  | 2.9667 | 2.45 | 1.5333 | 4.3333 | 3.1667
| Median |  | 3 | 2.3545 | 1.605 | 2.6 | 2 | 1.95 | 2 | 2 | 1.8 |  |  | 4 |  | 2.95 | 2.4 | 1.5 | 4 | 3
| Std Dev. |  | 0 | 0.13351 | 0.035071 | 0.083666 | 0.063246 | 0.054772 | 0.083666 | 0.05164 | 0.040825 |  |  |  |  | 0.17512 | 0.083666 | 0.10328 | 0.5164 | 0.40825
| Rel.Std.Dev. |  | 0% | 5.59% | 2.185% | 3.157% | 3.162% | 2.809% | 4.291% | 2.626% | 2.247% |  |  |  |  | 5.903% | 3.415% | 6.736% | 11.92% | 12.89%
| PDM3 |  | 47.71% | 17.59% | -20.97% | 30.48% | -1.526% | -3.988% | -3.988% | -3.168% | -10.55% |  |  | 96.95% |  | 46.07% | 20.63% | -24.5% | 113.4% | 55.92%
| Z-Score (Absolute) |  | NA | 0.71 | 0.56 | 1.13 | 0.08 | 0 | 0 | 0.03 | 0.22 | 2.35 | 0.89 | NA | 3.07 | NA | 0.81 | 0.67 | NA | NA
| Analytical results for Lu in OREAS 931b (Certified Value 0.092 ppm) | [24 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [5 empty] | 1
| No | No | 01: 03; 02: 10; 03: 12; 04: 15; 05: 17; 06: 20 | [5 empty] | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | [5 empty] | 3
|  |  | 01: 0.5g; 02: 0.25g; 03: 0.25g; 04: 01g; 05: 0.5g; 06: 0.25g | [5 empty] | 3
| 1 | 1 | 01: 0.077; 02: 0.09; 03: 0.11; 04: 0.1; 05: < 0.1; 06: 0.09 | [5 empty] | 1
| 1 | 2 | 01: 0.079; 02: 0.08; 03: 0.09; 04: 0.09; 05: < 0.1; 06: 0.09 | [5 empty] | 0
| 1 | 3 | 01: 0.076; 02: 0.09; 03: 0.11; 04: 0.1; 05: < 0.1; 06: 0.09 | [5 empty] | 16
| 1 | 4 | 01: 0.078; 02: 0.08; 03: 0.11; 04: 0.1; 05: < 0.1; 06: 0.09 | [5 empty] | 0.0915
| 1 | 5 | 01: 0.078; 02: 0.09; 03: 0.11; 04: 0.09; 05: < 0.1; 06: 0.09 | [5 empty] | 90
| 1 | 6 | 01: 0.077; 02: 0.09; 03: 0.12; 04: 0.09; 05: < 0.1; 06: 0.09
| Mean |  | 01: 0.0775; 02: 0.086667; 03: 0.10833; 04: 0.095; 06: 0.09
| Median |  | 01: 0.0775; 02: 0.09; 03: 0.11; 04: 0.095; 06: 0.09
| Std Dev. |  | 01: 0.0010488; 02: 0.005164; 03: 0.0098319; 04: 0.0054772; 06: 1.5202e-17
| Rel.Std.Dev. |  | 01: 1.353%; 02: 5.958%; 03: 9.076%; 04: 5.766%; 06: 1.689e-14%
| PDM3 |  | 01: -15.3%; 02: -5.282%; 03: 18.4%; 04: 3.825%; 06: -1.639%
| Z-Score (Absolute) |  | 01: 0.83; 02: 0.13; 03: 1.54; 04: 0.51; 05: 2.95; 06: 0.13
| Analytical results for Mg in OREAS 931b (Certified Value 0.128 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  | 3
| 1 | 1 | 0.15 | 0.13814 | 0.127 | 0.13 | 0.12 | 0.11 | 0.12 | 0.11 | 0.11 | 0.14 | 0.15815 | 0.135 | 0.13 | 0.12 | 0.14 | 0.12 | 0.14 | 0.13 | 0.1209 | 0.13 | 0.121 | 0.132 |  | 1
| 1 | 2 | 0.14 | 0.13854 | 0.127 | 0.12 | 0.12 | 0.12 | 0.12 | 0.11 | 0.11 | 0.14 | 0.14672 | 0.137 | 0.14 | 0.12 | 0.15 | 0.12 | 0.13 | 0.13 | 0.1207 | 0.12 | 0.121 | 0.131 |  | 0
| 1 | 3 | 0.15 | 0.13806 | 0.129 | 0.13 | 0.11 | 0.12 | 0.12 | 0.11 | 0.11 | 0.14 | 0.15488 | 0.137 | 0.15 | 0.12 | 0.14 | 0.12 | 0.12 | 0.13 | 0.1225 | 0.13 | 0.115 | 0.129 |  | 16
| 1 | 4 | 0.14 | 0.13218 | 0.126 | 0.13 | 0.12 | 0.12 | 0.12 | 0.11 | 0.11 | 0.14 | 0.15661 | 0.139 | 0.13 | 0.12 | 0.14 | 0.12 | 0.13 | 0.12 | 0.1239 | 0.13 | 0.115 | 0.131 |  | 0.12776
| 1 | 5 | 0.15 | 0.13302 | 0.126 | 0.13 | 0.11 | 0.12 | 0.12 | 0.11 | 0.11 | 0.14 | 0.14887 | 0.13 | 0.14 | 0.12 | 0.15 | 0.12 | 0.13 | 0.12 | 0.1169 | 0.12 | 0.121 | 0.135 |  | 91
| 1 | 6 | 0.15 | 0.13247 | 0.127 | 0.13 | 0.11 | 0.12 | 0.12 | 0.11 | 0.11 | 0.14 | 0.1407 | 0.142 | 0.14 | 0.12 | 0.15 | 0.12 | 0.13 | 0.12 | 0.119 | 0.12 | 0.115 | 0.131
| Mean |  | 0.14667 | 0.1354 | 0.127 | 0.12833 | 0.115 | 0.11833 | 0.12 | 0.11 | 0.11 | 0.14 | 0.15099 | 0.13667 | 0.13833 | 0.12 | 0.145 | 0.12 | 0.13 | 0.125 | 0.12065 | 0.125 | 0.118 | 0.1315
| Median |  | 0.15 | 0.13554 | 0.127 | 0.13 | 0.115 | 0.12 | 0.12 | 0.11 | 0.11 | 0.14 | 0.15187 | 0.137 | 0.14 | 0.12 | 0.145 | 0.12 | 0.13 | 0.125 | 0.1208 | 0.125 | 0.118 | 0.131
| Std Dev. |  | 0.005164 | 0.0031325 | 0.0010954 | 0.0040825 | 0.0054772 | 0.0040825 | 0 | 0 | 0 | 0 | 0.0067306 | 0.0040332 | 0.0075277 | 0 | 0.0054772 | 0 | 0.0063246 | 0.0054772 | 0.0024833 | 0.0054772 | 0.0032863 | 0.0019748
| Rel.Std.Dev. |  | 3.521% | 2.314% | 0.8626% | 3.181% | 4.763% | 3.45% | 0% | 0% | 0% | 0% | 4.458% | 2.951% | 5.442% | 0% | 3.777% | 0% | 4.865% | 4.382% | 2.058% | 4.382% | 2.785% | 1.502%
| PDM3 |  | 14.8% | 5.979% | -0.5966% | 0.447% | -9.989% | -7.38% | -6.076% | -13.9% | -13.9% | 9.579% | 18.18% | 6.97% | 8.274% | -6.076% | 13.49% | -6.076% | 1.752% | -2.162% | -5.567% | -2.162% | -7.641% | 2.926%
| Z-Score (Absolute) |  | 1.73 | 0.79 | 0.08 | 0.2 | 0.92 | 0.64 | 0.5 | 1.34 | 1.34 | 1.17 | 2.1 | 0.9 | 1.04 | 0.5 | 1.59 | 0.5 | 0.34 | 0.08 | 0.45 | 0.08 | 0.71 | 0.46
| Analytical results for Mn in OREAS 931b (Certified Value 0.014 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 3
| 1 | 1 | 0.0228 | 0.01367 | 0.01394 | 0.0127 | 0.0135 | 0.0139 | 0.0136 | 0.0129 | 0.013 | 0.0153 | 0.012454 | 0.0139 | 0.014 | 0.0116 | 0.0142 | 0.0139 | 0.0145 | 0.0142 | 0.016616 | 0.0144 | 0.0124 | 0.014 |  | 1
| 1 | 2 | 0.023 | 0.013448 | 0.01401 | 0.0124 | 0.0137 | 0.0141 | 0.0137 | 0.0133 | 0.013 | 0.015 | 0.012095 | 0.0141 | 0.0156 | 0.0122 | 0.0148 | 0.0146 | 0.014 | 0.014 | 0.016335 | 0.0141 | 0.0132 | 0.0141 |  | 0
| 1 | 3 | 0.0231 | 0.01349 | 0.01421 | 0.0126 | 0.0136 | 0.0143 | 0.0136 | 0.0131 | 0.0131 | 0.0151 | 0.012475 | 0.0141 | 0.0165 | 0.0129 | 0.0143 | 0.0138 | 0.0143 | 0.0139 | 0.016595 | 0.0142 | 0.0124 | 0.0139 |  | 16
| 1 | 4 | 0.0231 | 0.013659 | 0.01408 | 0.0131 | 0.0134 | 0.0142 | 0.0139 | 0.0133 | 0.0132 | 0.015 | 0.012475 | 0.0142 | 0.0144 | 0.0134 | 0.0143 | 0.014 | 0.0134 | 0.0137 | 0.016646 | 0.0142 | 0.0124 | 0.014 |  | 0.013706
| 1 | 5 | 0.023 | 0.013484 | 0.01401 | 0.0128 | 0.0136 | 0.0142 | 0.0137 | 0.0131 | 0.013 | 0.0152 | 0.01239 | 0.0135 | 0.0152 | 0.012 | 0.0144 | 0.0145 | 0.0141 | 0.0137 | 0.016427 | 0.0142 | 0.0132 | 0.0145 |  | 92
| 1 | 6 | 0.0229 | 0.013518 | 0.01404 | 0.0127 | 0.0135 | 0.0144 | 0.0132 | 0.0127 | 0.0124 | 0.0152 | 0.012412 | 0.0147 | 0.0156 | 0.0121 | 0.0148 | 0.0142 | 0.0139 | 0.0137 | 0.0161 | 0.0141 | 0.0124 | 0.0141
| Mean |  | 0.022983 | 0.013545 | 0.014048 | 0.012717 | 0.01355 | 0.014183 | 0.013617 | 0.013067 | 0.01295 | 0.015133 | 0.012383 | 0.014083 | 0.015217 | 0.012367 | 0.014467 | 0.014167 | 0.014033 | 0.013867 | 0.016453 | 0.0142 | 0.012667 | 0.0141
| Median |  | 0.023 | 0.013504 | 0.014025 | 0.0127 | 0.01355 | 0.0142 | 0.01365 | 0.0131 | 0.013 | 0.01515 | 0.012433 | 0.0141 | 0.0154 | 0.01215 | 0.01435 | 0.0141 | 0.01405 | 0.0138 | 0.016511 | 0.0142 | 0.0124 | 0.01405
| Std Dev. |  | 0.0001169 | 9.5373e-05 | 9.1524e-05 | 0.00023166 | 0.00010488 | 0.00017224 | 0.00023166 | 0.00023381 | 0.00028107 | 0.00012111 | 0.00014552 | 0.000392 | 0.00090425 | 0.00065929 | 0.00026583 | 0.0003266 | 0.00037771 | 0.00020656 | 0.00021138 | 0.00010954 | 0.00041312 | 0.00020976
| Rel.Std.Dev. |  | 0.5086% | 0.7041% | 0.6515% | 1.822% | 0.774% | 1.214% | 1.701% | 1.789% | 2.17% | 0.8003% | 1.175% | 2.783% | 5.942% | 5.331% | 1.838% | 2.305% | 2.692% | 1.49% | 1.285% | 0.7714% | 3.261% | 1.488%
| PDM3 |  | 67.69% | -1.174% | 2.499% | -7.217% | -1.137% | 3.484% | -0.6504% | -4.663% | -5.515% | 10.42% | -9.648% | 2.754% | 11.02% | -9.771% | 5.551% | 3.362% | 2.39% | 1.174% | 20.04% | 3.606% | -7.582% | 2.876%
| Z-Score (Absolute) |  | 12.22 | 0.68 | 0.01 | 1.81 | 0.67 | 0.19 | 0.58 | 1.33 | 1.49 | 1.49 | 2.27 | 0.06 | 1.61 | 2.29 | 0.58 | 0.17 | 0.01 | 0.24 | 3.3 | 0.22 | 1.9 | 0.08
| Analytical results for Mo in OREAS 931b (Certified Value 3.4 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 3
| 1 | 1 | 3 | 3.0122 | 3.4 | 3.2 | 3.27 | 3.4 | 3.3 | 3.22 | 3.17 | 3.84 | < 5 | 3.48 | 3.86 | 3.53 | 3.3 | 3 | 3.47 | 3.46 | 3.4497 | 3.61 | 2 | 3 | 3.1376 | 1
| 1 | 2 | 3 | 3.115 | 3.36 | 3.3 | 3.39 | 3.5 | 3.4 | 3.26 | 3.28 | 3.89 | < 5 | 3.54 | 4.38 | 3.56 | 3.4 | 3.1 | 3.4 | 3.63 | 3.4423 | 3.43 | 2 | 4 | 3.3985 | 8
| 1 | 3 | 3 | 3.4843 | 3.42 | 3.3 | 3.14 | 3.67 | 3.23 | 3.19 | 3.29 | 3.96 | < 5 | 3.72 | 4.63 | 3.47 | 3.4 | 2.9 | 3.47 | 3.61 | 3.9546 | 3.69 | 2 | 4 | 3.3681 | 16
| 1 | 4 | 3 | 3.4403 | 3.39 | 3.2 | 3.27 | 3.64 | 3.29 | 3.16 | 3.35 | 3.94 | < 5 | 3.78 | 4.12 | 3.52 | 3.4 | 2.9 | 3.39 | 3.62 | 4.8554 | 3.66 | 2 | 4 | 3.3257 | 3.4024
| 1 | 5 | 3 | 3.2197 | 3.36 | 3.2 | 3.2 | 3.73 | 3.34 | 3.1 | 3.23 | 3.92 | < 5 | 3.63 | 4.17 | 3.55 | 3.3 | 3 | 3.5 | 3.5 | 4.5183 | 3.62 | 2 | 4 | 3.4978 | 93
| 1 | 6 | 3 | 3.0859 | 3.38 | 3.3 | 3.17 | 3.62 | 3.07 | 3.09 | 3.19 | 3.9 | < 5 | 3.86 | 4.61 | 3.52 | 3.2 | 3 | 3.4 | 3.48 | 4.091 | 3.69 | 2 | 3 | 3.3712
| Mean |  | 3 | 3.2263 | 3.385 | 3.25 | 3.24 | 3.5933 | 3.2717 | 3.17 | 3.2517 | 3.9083 |  | 3.6683 | 4.295 | 3.525 | 3.3333 | 2.9833 | 3.4383 | 3.55 | 4.0519 | 3.6167 | 2 | 3.6667 | 3.3498
| Median |  | 3 | 3.1674 | 3.385 | 3.25 | 3.235 | 3.63 | 3.295 | 3.175 | 3.255 | 3.91 |  | 3.675 | 4.275 | 3.525 | 3.35 | 3 | 3.435 | 3.555 | 4.0228 | 3.64 | 2 | 4 | 3.3697
| Std Dev. |  | 0 | 0.19513 | 0.023452 | 0.054772 | 0.090333 | 0.12127 | 0.11374 | 0.066933 | 0.067651 | 0.042151 |  | 0.14511 | 0.30138 | 0.031464 | 0.08165 | 0.075277 | 0.047081 | 0.077974 | 0.56718 | 0.097502 | 0 | 0.5164 | 0.11889
| Rel.Std.Dev. |  | 0% | 6.048% | 0.6928% | 1.685% | 2.788% | 3.375% | 3.476% | 2.111% | 2.081% | 1.078% |  | 3.956% | 7.017% | 0.8926% | 2.449% | 2.523% | 1.369% | 2.196% | 14% | 2.696% | 0% | 14.08% | 3.549%
| PDM3 |  | -11.83% | -5.177% | -0.5115% | -4.479% | -4.773% | 5.612% | -3.842% | -6.831% | -4.43% | 14.87% |  | 7.816% | 26.23% | 3.603% | -2.03% | -12.32% | 1.056% | 4.338% | 19.09% | 6.297% | -41.22% | 7.767% | -1.546%
| Z-Score (Absolute) |  | NA | 0.56 | 0.07 | 0.47 | 0.51 | 0.9 | 0.38 | 0.78 | 0.46 | 2.14 | 3.44 | 1.19 | 3.68 | 0.62 | 0.14 | 1.52 | 0.28 | 0.72 | 2.71 | 0.99 | NA | NA | 0.07
Results from laboratories 1, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 931b (Certified Value 0.01 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  | 3
| 1 | 1 | 0.01 | 0.010936 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.020654 | < 0.005 | 0.009 | < 0.01 | 0.01 | 0.007 | 0.014 | 0.01 | 0.1209 | < 0.01 | < 0.007 | 0.015 |  | 1
| 1 | 2 | 0.01 | 0.018914 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.019737 | < 0.005 | 0.01 | < 0.01 | 0.01 | 0.008 | 0.013 | 0.01 | 0.1207 | < 0.01 | < 0.007 | 0.013 |  | 19
| 1 | 3 | 0.01 | 0.011721 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.02 | < 0.01 | 0.020439 | < 0.005 | 0.011 | < 0.01 | 0.01 | 0.007 | 0.013 | 0.01 | 0.1225 | < 0.01 | < 0.007 | 0.015 |  | 16
| 1 | 4 | 0.01 | 0.019163 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.020666 | < 0.005 | 0.009 | < 0.01 | 0.01 | 0.007 | 0.013 | 0.01 | 0.1239 | < 0.01 | < 0.007 | 0.013 |  | 0.010141
| 1 | 5 | 0.01 | 0.010294 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.019977 | < 0.005 | 0.01 | < 0.01 | 0.01 | 0.008 | 0.013 | 0.01 | 0.1169 | < 0.01 | < 0.007 | 0.014 |  | 94
| 1 | 6 | 0.01 | 0.020096 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | < 0.01 | 0.01 | 0.01 | 0.020477 | < 0.005 | 0.01 | < 0.01 | 0.01 | 0.007 | 0.013 | 0.01 | 0.119 | < 0.01 | < 0.007 | 0.011
| Mean |  | 0.01 | 0.015188 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.018333 | 0.01 | 0.020325 |  | 0.0098333 |  | 0.01 | 0.0073333 | 0.013167 | 0.01 | 0.12065 |  |  | 0.0135
| Median |  | 0.01 | 0.015317 | 0.008 | 0.01 | 0.02 | 0.01 | 0.01 | 0.01 | 0.02 | 0.01 | 0.020458 |  | 0.01 |  | 0.01 | 0.007 | 0.013 | 0.01 | 0.1208 |  |  | 0.0135
| Std Dev. |  | 0 | 0.0046437 | 0 | 0 | 0 | 0 | 0 | 0 | 0.0040825 | 0 | 0.00038152 |  | 0.00075277 |  | 0 | 0.0005164 | 0.00040825 | 0 | 0.0024833 |  |  | 0.0015166
| Rel.Std.Dev. |  | 0% | 30.58% | 0% | 0% | 0% | 0% | 0% | 0% | 22.27% | 0% | 1.877% |  | 7.655% |  | 0% | 7.042% | 3.101% | 0% | 2.058% |  |  | 11.23%
| PDM3 |  | -1.391% | 49.76% | -21.11% | -1.391% | 97.22% | -1.391% | -1.391% | -1.391% | 80.78% | -1.391% | 100.4% |  | -3.034% |  | -1.391% | -27.69% | 29.84% | -1.391% | 1090% |  |  | 33.12%
| Z-Score (Absolute) |  | 0 | 1.2 | 0.46 | 0 | 2.31 | 0 | 0 | 0.19 | 1.93 | 0.19 | 2.39 | 1.73 | 0.04 | 1.16 | 0 | 0.62 | 0.73 | 0 | 25.58 | 1.16 | 1.45 | 0.81
| Analytical results for Nb in OREAS 931b (Certified Value 0.056 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | [6 empty] | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | [6 empty] | 3
| 1 | 1 | < 5 | < 0.2 | < 0.05 | < 0.2 | 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | 0.15 | 0.11 | 0.06 | < 0.1 | < 0.05 | < 0.1 | 0.08 | [6 empty] | 1
| 1 | 2 | < 5 | < 0.2 | < 0.05 | < 0.2 | 0.05 | 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.19 | 0.1 | 0.06 | < 0.1 | < 0.05 | < 0.1 | 0.07 | [6 empty] | 10
| 1 | 3 | < 5 | < 0.2 | < 0.05 | < 0.2 | 0.05 | 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.13 | 0.1 | 0.06 | < 0.1 | < 0.05 | < 0.1 | 0.07 | [6 empty] | 16
| 1 | 4 | < 5 | < 0.2 | < 0.05 | < 0.2 | 0.05 | 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.13 | 0.1 | 0.05 | < 0.1 | < 0.05 | < 0.1 | 0.07 | [6 empty] | 0.055556
| 1 | 5 | < 5 | < 0.2 | < 0.05 | < 0.2 | 0.05 | 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | 0.13 | 0.14 | 0.04 | 0.1 | < 0.05 | < 0.1 | 0.06 | [6 empty] | 95
| 1 | 6 | < 5 | < 0.2 | < 0.05 | < 0.2 | 0.05 | 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.05 | 0.13 | 0.11 | 0.05 | < 0.1 | < 0.05 | < 0.1 | 0.06
| Mean |  | 05: 0.05; 06: 0.051667; 07: 0.06; 08: 0.05; 12: 0.14333; 14: 0.11; 15: 0.053333; 17: 0.1; 20: 0.068333
| Median |  | 05: 0.05; 06: 0.05; 07: 0.06; 08: 0.05; 12: 0.13; 14: 0.105; 15: 0.055; 17: 0.1; 20: 0.07
| Std Dev. |  | 05: 7.6012e-18; 06: 0.0040825; 08: 0; 12: 0.024221; 14: 0.015492; 15: 0.008165; 20: 0.0075277
| Rel.Std.Dev. |  | 05: 1.52e-14%; 06: 7.902%; 08: 0%; 12: 16.9%; 14: 14.08%; 15: 15.31%; 20: 11.02%
| PDM3 |  | 05: -10%; 06: -7%; 07: 8%; 08: -10%; 12: 158%; 14: 98%; 15: -4%; 17: 80%; 20: 23%
| Z-Score (Absolute) |  | 61.91 | 1.22 | 0.67 | 1.22 | 0.04 | 0 | 0.53 | 0.25 | 0.67 | 0.67 | 2.32 | 1.48 | 0.04 | 0.17 | 0.67 | 0.04 | 0.42
| Analytical results for Nd in OREAS 931b (Indicative Value 17 ppm) | [24 empty] | Indicative
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | [6 empty] | 1
| No | No | 01: 03; 02: 15; 03: 17 | [6 empty] | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS | [6 empty] | 1
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g | [6 empty] | 1
| 1 | 1 | 01: 16.546; 02: 26.36; 03: 17.7 | [6 empty] | 1
| 1 | 2 | 01: 16.952; 02: 25.97; 03: 17.3 | [6 empty] | 5
| 1 | 3 | 01: 16.821; 02: 27.24; 03: 17.5 | [6 empty] | 16
| 1 | 4 | 01: 16.637; 02: 26.18; 03: 16.7 | [6 empty] | 16.9967
| 1 | 5 | 01: 16.427; 02: 24.47; 03: 17.6 | [6 empty] | 11
| 1 | 6 | 01: 17.077; 02: 25.81; 03: 16.7
| Mean |  | 01: 16.7433; 02: 26.005; 03: 17.25
| Median |  | 01: 16.729; 02: 26.075; 03: 17.4
| Std Dev. |  | 01: 0.24949; 02: 0.90303; 03: 0.44609
| Rel.Std.Dev. |  | 01: 1.49%; 02: 3.473%; 03: 2.586%
| PDM3 |  | 01: -1.49%; 02: 53%; 03: 1.49%
| Z-Score (Absolute) |  | 01: 0.67; 02: 11.65; 03: 0
| Analytical results for Ni in OREAS 931b (Certified Value 10.2 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 2
| 1 | 1 | 10 | 9.8076 | 10.6 | 9 | 11 | 9.9 | 10.9 | 9.4 | 9.7 | 12.2 | 9.389 | 5.8 | 8.8 | 10.4 | 11.1 | 9.9 | 9.9 | 10 | 8.6961 | 11.4 | 10 | 11 |  | 1
| 1 | 2 | 10 | 9.9386 | 10.6 | 9 | 10.7 | 9.9 | 11 | 9.6 | 10.6 | 11.2 | 9.285 | 4.1 | 10 | 10.3 | 10.6 | 10.1 | 9.2 | 10.4 | 8.9936 | 11 | 9 | 12 |  | 12
| 1 | 3 | 10 | 10.3 | 10.6 | 10 | 10.7 | 10.1 | 10.6 | 9.3 | 10.2 | 12.2 | 9.767 | 6 | 10.5 | 10.5 | 11.2 | 9.7 | 9.7 | 10.3 | 9.3051 | 11.6 | 9 | 12 |  | 16
| 1 | 4 | 10 | 10.5409 | 10.4 | 9 | 10.7 | 10.1 | 10.5 | 9.5 | 10 | 11.7 | 9.146 | 5.9 | 9 | 10.5 | 10.9 | 9.6 | 8.9 | 10.1 | 9.1331 | 11.3 | 9 | 10 |  | 10.2019
| 1 | 5 | 10 | 10.4188 | 10.7 | 9 | 10.7 | 10 | 10.8 | 9 | 10.4 | 11.8 | 9.391 | 5.7 | 9.7 | 10.6 | 10.4 | 9.8 | 9.5 | 9.9 | 9.0755 | 11.3 | 9 | 11 |  | 96
| 1 | 6 | 10 | 9.9773 | 10.6 | 9 | 10.5 | 10.2 | 10.2 | 9.4 | 10.2 | 11.8 | 9.697 | 5.9 | 9.8 | 10 | 10.5 | 9.7 | 9.1 | 10 | 8.9309 | 11.4 | 9 | 10
| Mean |  | 10 | 10.1639 | 10.5833 | 9.1667 | 10.7167 | 10.0333 | 10.6667 | 9.3667 | 10.1833 | 11.8167 | 9.4458 | 5.5667 | 9.6333 | 10.3833 | 10.7833 | 9.8 | 9.3833 | 10.1167 | 9.0224 | 11.3333 | 9.1667 | 11
| Median |  | 10 | 10.1387 | 10.6 | 9 | 10.7 | 10.05 | 10.7 | 9.4 | 10.2 | 11.8 | 9.39 | 5.85 | 9.75 | 10.45 | 10.75 | 9.75 | 9.35 | 10.05 | 9.0346 | 11.35 | 9 | 11
| Std Dev. |  | 0 | 0.29603 | 0.098319 | 0.40825 | 0.16021 | 0.12111 | 0.29439 | 0.20656 | 0.31252 | 0.37103 | 0.2401 | 0.72572 | 0.63456 | 0.2137 | 0.33116 | 0.17889 | 0.38166 | 0.19408 | 0.20524 | 0.19664 | 0.40825 | 0.89443
| Rel.Std.Dev. |  | 0% | 2.913% | 0.929% | 4.454% | 1.495% | 1.207% | 2.76% | 2.205% | 3.069% | 3.14% | 2.542% | 13.04% | 6.587% | 2.058% | 3.071% | 1.825% | 4.067% | 1.918% | 2.275% | 1.735% | 4.454% | 8.131%
| PDM3 |  | -1.979% | -0.3725% | 3.739% | -10.15% | 5.046% | -1.652% | 4.556% | -8.187% | -0.1819% | 15.83% | -7.411% | -45.43% | -5.573% | 1.779% | 5.699% | -3.939% | -8.024% | -0.8354% | -11.56% | 11.09% | -10.15% | 7.823%
| Z-Score (Absolute) |  | NA | 0.03 | 0.54 | NA | 0.7 | 0.13 | 0.64 | 0.95 | 0.05 | 2.05 | 0.85 | 5.59 | 0.62 | 0.3 | 0.79 | 0.42 | 0.93 | 0.03 | 1.37 | 1.46 | NA | NA
Results from laboratories 1, 4, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 931b (Certified Value 0.022 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 23 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.029 | 0.022276 | 0.0205 | 0.0231 | 0.022 | 0.02 | 0.022 | 0.021 | 0.019 | 0.0241 | 0.030123 | 0.019 | 0.0226 | 0.03 | 0.0231 | 0.019 | 0.035 | 0.0209 | 0.02 | < 0.0009 | 0.023 |  |  | 1
| 1 | 2 | 0.0288 | 0.022647 | 0.0192 | 0.0237 | 0.022 | 0.021 | 0.022 | 0.021 | 0.019 | 0.0237 | 0.029436 | 0.02 | 0.0255 | 0.031 | 0.0236 | 0.02 | 0.033 | 0.0208 | 0.02 | < 0.0009 | 0.02 |  |  | 0
| 1 | 3 | 0.0287 | 0.022509 | 0.0195 | 0.0232 | 0.022 | 0.02 | 0.021 | 0.021 | 0.019 | 0.0238 | 0.03046 | 0.02 | 0.0267 | 0.031 | 0.0226 | 0.018 | 0.034 | 0.0204 | 0.02 | < 0.0009 | 0.021 |  |  | 16
| 1 | 4 | 0.0286 | 0.021398 | 0.0195 | 0.0223 | 0.022 | 0.021 | 0.022 | 0.021 | 0.019 | 0.0235 | 0.030732 | 0.02 | 0.0233 | 0.033 | 0.0231 | 0.018 | 0.034 | 0.0201 | 0.02 | < 0.0009 | 0.022 |  |  | 0.021728
| 1 | 5 | 0.0288 | 0.021898 | 0.0192 | 0.023 | 0.022 | 0.021 | 0.022 | 0.021 | 0.018 | 0.0242 | 0.030311 | 0.019 | 0.0248 | 0.032 | 0.0231 | 0.019 | 0.034 | 0.0201 | 0.02 | < 0.0009 | 0.022 |  |  | 97
| 1 | 6 | 0.0284 | 0.021777 | 0.0192 | 0.023 | 0.022 | 0.021 | 0.021 | 0.021 | 0.018 | 0.0244 | 0.030288 | 0.02 | 0.0253 | 0.031 | 0.0233 | 0.019 | 0.035 | 0.0202 | 0.02 | < 0.0009 | 0.021
| Mean |  | 0.028717 | 0.022084 | 0.019517 | 0.02305 | 0.022 | 0.020667 | 0.021667 | 0.021 | 0.018667 | 0.02395 | 0.030225 | 0.019667 | 0.0247 | 0.031333 | 0.023133 | 0.018833 | 0.034167 | 0.020417 | 0.02 |  | 0.0215
| Median |  | 0.02875 | 0.022087 | 0.01935 | 0.02305 | 0.022 | 0.021 | 0.022 | 0.021 | 0.019 | 0.02395 | 0.030299 | 0.02 | 0.02505 | 0.031 | 0.0231 | 0.019 | 0.034 | 0.0203 | 0.02 |  | 0.0215
| Std Dev. |  | 0.00020412 | 0.00047613 | 0.00050365 | 0.00045056 | 0 | 0.0005164 | 0.0005164 | 0 | 0.0005164 | 0.00033912 | 0.00043714 | 0.0005164 | 0.0015086 | 0.0010328 | 0.0003266 | 0.00075277 | 0.00075277 | 0.00035449 | 0 |  | 0.0010488
| Rel.Std.Dev. |  | 0.7108% | 2.156% | 2.581% | 1.955% | 0% | 2.499% | 2.383% | 0% | 2.766% | 1.416% | 1.446% | 2.626% | 6.108% | 3.296% | 1.412% | 3.997% | 2.203% | 1.736% | 0% |  | 4.878%
| PDM3 |  | 32.17% | 1.641% | -10.18% | 6.086% | 1.253% | -4.883% | -0.2809% | -3.349% | -14.09% | 10.23% | 39.11% | -9.486% | 13.68% | 44.21% | 6.469% | -13.32% | 57.25% | -6.034% | -7.952% |  | -1.048%
| Z-Score (Absolute) |  | 2.38 | 0.14 | 0.72 | 0.47 | 0.11 | 0.34 | 0 | 0.22 | 1.01 | 0.77 | 2.89 | 0.67 | 1.02 | 3.26 | 0.49 | 0.96 | 4.21 | 0.42 | 0.56 | 7.16 | 0.06
| Analytical results for Pb in OREAS 931b (Certified Value 115 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | 0
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 0
| 1 | 1 | 115 | 119.1042 | 114.2 | 117 | 113.5 | 109.5 | 118 | 110.5 | 113 | 121.4 | 104.099 | 102 | 109.75 | 122 | 112 | 111.5 | 114 | 119.5 | 127.7492 | 115 | 203 | 126 | 114.6969 | 1
| 1 | 2 | 114 | 109.7265 | 114.5 | 115 | 112.5 | 110.5 | 118 | 110.5 | 113 | 119.4 | 102.957 | 98.8 | 123.42 | 122 | 115 | 114.7 | 113 | 119 | 131.1963 | 112 | 203 | 119 | 111.6861 | 13
| 1 | 3 | 113 | 113.1886 | 115.2 | 115 | 114.5 | 111.5 | 118.5 | 109.5 | 114.5 | 122 | 104.783 | 106 | 129.87 | 122 | 112 | 108.4 | 116 | 118.8 | 127.426 | 114 | 189 | 122 | 118.0932 | 16
| 1 | 4 | 116 | 110.5692 | 117.6 | 118 | 114 | 109.5 | 118.5 | 110.5 | 113.5 | 118.7 | 107.889 | 106 | 112.35 | 125 | 113 | 106.9 | 113 | 116.4 | 128.832 | 117 | 204 | 118 | 114.9514 | 115.1026
| 1 | 5 | 114 | 109.8188 | 114.4 | 115 | 112.5 | 110.5 | 117 | 112.5 | 112.5 | 121.5 | 105.544 | 99.7 | 119.56 | 122 | 112 | 108.8 | 119 | 118.6 | 127.7271 | 112 | 201 | 126 | 116.7851 | 98
| 1 | 6 | 114 | 117.9595 | 117.1 | 117 | 113.5 | 112.5 | 114 | 109.5 | 110 | 123.3 | 105.079 | 109 | 121.54 | 120 | 117 | 109.6 | 114 | 118.5 | 123.9595 | 117 | 185 | 124 | 116.7361
| Mean |  | 114.3333 | 113.3945 | 115.5 | 116.1667 | 113.4167 | 110.6667 | 117.3333 | 110.5 | 112.75 | 121.05 | 105.0585 | 103.5833 | 119.415 | 122.1667 | 113.5 | 109.9833 | 114.8333 | 118.4667 | 127.815 | 114.5 | 197.5 | 122.5 | 115.4915
| Median |  | 114 | 111.8789 | 114.85 | 116 | 113.5 | 110.5 | 118 | 110.5 | 113 | 121.45 | 104.931 | 104 | 120.55 | 122 | 112.5 | 109.2 | 114 | 118.7 | 127.7381 | 114.5 | 202 | 123 | 115.8438
| Std Dev. |  | 1.0328 | 4.1882 | 1.4805 | 1.3292 | 0.80104 | 1.169 | 1.7224 | 1.0954 | 1.5083 | 1.705 | 1.6529 | 4.037 | 7.392 | 1.6021 | 2.0736 | 2.7607 | 2.3166 | 1.0727 | 2.3451 | 2.2583 | 8.2885 | 3.4496 | 2.2537
| Rel.Std.Dev. |  | 0.9033% | 3.693% | 1.282% | 1.144% | 0.7063% | 1.056% | 1.468% | 0.9914% | 1.338% | 1.409% | 1.573% | 3.897% | 6.19% | 1.311% | 1.827% | 2.51% | 2.017% | 0.9055% | 1.835% | 1.972% | 4.197% | 2.816% | 1.951%
| PDM3 |  | -0.6684% | -1.484% | 0.3452% | 0.9244% | -1.465% | -3.854% | 1.938% | -3.999% | -2.044% | 5.167% | -8.726% | -10.01% | 3.747% | 6.137% | -1.392% | -4.448% | -0.234% | 2.923% | 11.04% | -0.5236% | 71.59% | 6.427% | 0.3378%
| Z-Score (Absolute) |  | 0.09 | 0.27 | 0.12 | 0.25 | 0.26 | 0.77 | 0.46 | 0.8 | 0.39 | 1.15 | 1.81 | 2.09 | 0.85 | 1.36 | 0.25 | 0.9 | 0 | 0.67 | 2.41 | 0.06 | 15.34 | 1.42 | 0.12
| Analytical results for Pd in OREAS 931b (Indicative Value 150 ppb) | [24 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab | 1
| No | No | 01: 02; 02: 03; 03: 14; 04: 15 | ppb
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS | 0
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 01g | 0
| 1 | 1 | 01: < 10; 02: 4; 03: 304; 04: < 10 | 1
| 1 | 2 | 01: < 10; 02: 4; 03: 282; 04: < 10 | 6
| 1 | 3 | 01: < 10; 02: 4; 03: 301; 04: < 10 | 16
| 1 | 4 | 01: < 10; 02: 4; 03: 298; 04: < 10 | 149.8333
| 1 | 5 | 01: < 10; 02: 4; 03: 296; 04: < 10 | 12
| 1 | 6 | 01: < 10; 02: 4; 03: 293; 04: < 10
| Mean |  | 02: 4; 03: 295.6667
| Median |  | 02: 4; 03: 297
| Std Dev. |  | 02: 0; 03: 7.7115
| Rel.Std.Dev. |  | 02: 0%; 03: 2.608%
| PDM3 |  | 02: -97.33%; 03: 97.33%
| Z-Score (Absolute) |  | 01: 0; 02: 1.35; 03: 392; 04: 0
| Analytical results for Pr in OREAS 931b (Indicative Value 5.43 ppm) | [24 empty] | Indicative
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab | 1
| No | No | 01: 03; 02: 15; 03: 17 | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g | 2
| 1 | 1 | 01: 4.37; 02: 7.2; 03: 4.9 | 1
| 1 | 2 | 01: 4.437; 02: 7.05; 03: 4.8 | 14
| 1 | 3 | 01: 4.399; 02: 7.41; 03: 4.9 | 16
| 1 | 4 | 01: 4.405; 02: 7; 03: 4.7 | 5.4337
| 1 | 5 | 01: 4.341; 02: 6.67; 03: 5 | 9
| 1 | 6 | 01: 4.495; 02: 7.03; 03: 4.7
| Mean |  | 01: 4.4078; 02: 7.06; 03: 4.8333
| Median |  | 01: 4.402; 02: 7.04; 03: 4.85
| Std Dev. |  | 01: 0.053727; 02: 0.24446; 03: 0.12111
| Rel.Std.Dev. |  | 01: 1.219%; 02: 3.463%; 03: 2.506%
| PDM3 |  | 01: -18.88%; 02: 29.93%; 03: -11.05%
| Z-Score (Absolute) |  | 01: 0.67; 02: 3.53; 03: 0
| Analytical results for Pt in OREAS 931b (Indicative Value < 5 ppb) | [24 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab | 1
| No | No | 01: 02; 02: 03; 03: 14; 04: 15 | ppb
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS | 2
|  |  | 01: 01g; 02: 0.5g; 03: 0.5g; 04: 01g | 2
| 1 | 1 | 01: < 5; 02: 2; 03: < 10; 04: < 5 | 1
| 1 | 2 | 01: < 5; 02: 2; 03: < 10; 04: < 5 | 4
| 1 | 3 | 01: < 5; 02: 2; 03: < 10; 04: < 5 | 16
| 1 | 4 | 01: < 5; 02: 2; 03: < 10; 04: < 5 | < 5
| 1 | 5 | 01: < 5; 02: 2; 03: < 10; 04: < 5 | 10
| 1 | 6 | 01: < 5; 02: 2; 03: < 10; 04: < 5
| Mean |  | 02: 2
| Median |  | 02: 2
| Std Dev. |  | 02: 0
Rel.Std.Dev. | bar: 02
PDM3
| Z-Score (Absolute) |  | 01: 0; 02: 1.35; 03: 6.74; 04: 0
| Analytical results for Rb in OREAS 931b (Certified Value 13.7 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [6 empty] | 1
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | [6 empty] | 2
| 1 | 1 | 14.8409 | 13.6 | 20.4 | 14.3 | 12.7 | 14.4 | 12 | 12.3 | 14.9 | 16.264 | 12.58 | 12.3 | 23.83 | 15.2 | 14.5 | 20.8029 | 13.4 | [6 empty] | 1
| 1 | 2 | 14.6547 | 13.63 | 20.8 | 13.6 | 12.6 | 14.6 | 12 | 13.5 | 14.2 | 15.765 | 12.57 | 12.4 | 23.28 | 14.4 | 14.6 | 21.7327 | 13.5 | [6 empty] | 15
| 1 | 3 | 14.8465 | 14.12 | 20.2 | 13.3 | 12.9 | 13.8 | 12.2 | 12.9 | 13.8 | 17.037 | 13.19 | 12.3 | 24.4 | 14.1 | 15 | 21.713 | 14.3 | [6 empty] | 16
| 1 | 4 | 14.167 | 14.08 | 20.6 | 13.8 | 12.7 | 13.8 | 12.2 | 12.6 | 14.2 | 17.165 | 13.07 | 12.8 | 22.41 | 13.9 | 14.2 | 22.9229 | 14.9 | [6 empty] | 13.7467
| 1 | 5 | 14.3988 | 13.93 | 20.3 | 13.6 | 13.1 | 14 | 12.2 | 12.8 | 14.1 | 15.447 | 12.67 | 12.6 | 21 | 14.5 | 13.9 | 20.5375 | 14.6 | [6 empty] | 99
| 1 | 6 | 14.4617 | 14.22 | 20.2 | 13.9 | 13.3 | 13.6 | 11.8 | 12 | 14.2 | 15.248 | 13.48 | 12 | 22.57 | 14.7 | 14.5 | 18.3439 | 14.8
| Mean |  | 14.5616 | 13.93 | 20.4167 | 13.75 | 12.8833 | 14.0333 | 12.0667 | 12.6833 | 14.2333 | 16.1543 | 12.9267 | 12.4 | 22.915 | 14.4667 | 14.45 | 21.0088 | 14.25
| Median |  | 14.5582 | 14.005 | 20.35 | 13.7 | 12.8 | 13.9 | 12.1 | 12.7 | 14.2 | 16.0145 | 12.87 | 12.35 | 22.925 | 14.45 | 14.5 | 21.258 | 14.45
| Std Dev. |  | 0.26834 | 0.26138 | 0.24014 | 0.33912 | 0.27142 | 0.38816 | 0.1633 | 0.51929 | 0.36148 | 0.81067 | 0.37665 | 0.27568 | 1.2014 | 0.45898 | 0.37283 | 1.5524 | 0.65345
| Rel.Std.Dev. |  | 1.843% | 1.876% | 1.176% | 2.466% | 2.107% | 2.766% | 1.353% | 4.094% | 2.54% | 5.018% | 2.914% | 2.223% | 5.243% | 3.173% | 2.58% | 7.389% | 4.586%
| PDM3 |  | 5.928% | 1.333% | 48.52% | 0.02401% | -6.281% | 2.085% | -12.22% | -7.735% | 3.54% | 17.51% | -5.965% | -9.797% | 66.69% | 5.237% | 5.116% | 52.83% | 3.661%
| Z-Score (Absolute) |  | 0.17 | 0.16 | 3.19 | 0.25 | 0.7 | 0.1 | 1.12 | 0.8 | 0 | 0.99 | 0.67 | 0.95 | 4.48 | 0.12 | 0.11 | 3.5 | 0.01
| Analytical results for Re in OREAS 931b (Certified Value 0.001 ppm) | [24 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 14 | 15 | 17 | 18 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | < 0.05 | 0.001 | < 0.05 | 0.001 | 0.002 | 0.001 | 0.002 | 0.002 | < 0.1 | < 0.001 | < 0.001 | 0.001 | 0.001 | [10 empty] | 1
| 1 | 2 | < 0.05 | 0.001 | < 0.05 | 0.002 | 0.002 | 0.001 | 0.001 | 0.001 | < 0.1 | < 0.001 | 0.001 | 0.001 | 0.001 | [10 empty] | 0
| 1 | 3 | < 0.05 | 0.001 | < 0.05 | 0.002 | 0.001 | 0.001 | 0.001 | 0.001 | < 0.1 | < 0.001 | 0.002 | 0.001 | 0.002 | [10 empty] | 16
| 1 | 4 | < 0.05 | 0.001 | < 0.05 | 0.002 | 0.001 | 0.002 | 0.002 | 0.001 | < 0.1 | < 0.001 | 0.001 | 0.001 | 0.001 | [10 empty] | 0.0012813
| 1 | 5 | < 0.05 | 0.001 | < 0.05 | 0.002 | 0.002 | 0.001 | 0.001 | 0.002 | < 0.1 | < 0.001 | 0.001 | 0.002 | 0.001 | [10 empty] | 100
| 1 | 6 | < 0.05 | 0.001 | < 0.05 | 0.002 | 0.002 | 0.002 | 0.001 | 0.001 | < 0.1 | < 0.001 | 0.003 | 0.001 | 0.001
| Mean |  |  | 0.001 |  | 0.0018333 | 0.0016667 | 0.0013333 | 0.0013333 | 0.0013333 |  |  | 0.0016 | 0.0011667 | 0.0011667
| Median |  |  | 0.001 |  | 0.002 | 0.002 | 0.001 | 0.001 | 0.001 |  |  | 0.001 | 0.001 | 0.001
| Std Dev. |  |  | 0 |  | 0.00040825 | 0.0005164 | 0.0005164 | 0.0005164 | 0.0005164 |  |  | 0.00089443 | 0.00040825 | 0.00040825
| Rel.Std.Dev. |  |  | 0% |  | 22.27% | 30.98% | 38.73% | 38.73% | 38.73% |  |  | 55.9% | 34.99% | 34.99%
| PDM3 |  |  | -21.95% |  | 43.09% | 30.08% | 4.065% | 4.065% | 4.065% |  |  | 24.88% | -8.943% | -8.943%
| Z-Score (Absolute) |  | 47.88 | 0.67 | 47.88 | 1.01 | 0.67 | 0 | 0 | 0 | 98.45 | 1.69 | 0.17 | 0.34 | 0.34
| Analytical results for S in OREAS 931b (Certified Value 4.25 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g | 0.5g |  | 3
| 1 | 1 | 4.18 | 4.4425 | 4.04 | 4.25 | 4.23 | 4.36 | 4.39 | 4.05 | 4.09 | 4.94 | 4.35 | 3.95 | 3.96 | 4.3012 | 4.02 | 4.352 | 4.65 | 3.991 | 4.38 | 3.06 | 4.34 | 3.8912 |  | 1
| 1 | 2 | 4.14 | 4.5356 | 3.8 | 4.18 | 4.28 | 4.41 | 4.41 | 4.09 | 4.1 | 4.92 | 4.4 | 4.36 | 3.84 | 4.557 | 4.16 | 4.236 | 4.63 | 4.464 | 4.33 | 3.06 | 4.33 | 3.9009 |  | 0
| 1 | 3 | 4.25 | 4.5272 | 3.9 | 4.23 | 4.26 | 4.43 | 4.4 | 4.08 | 4.1 | 4.98 | 4.37 | 4.63 | 3.88 | 4.3551 | 4.02 | 4.326 | 4.62 | 4.713 | 4.33 | 2.97 | 4.31 | 3.9508 |  | 16
| 1 | 4 | 4.16 | 4.4318 | 3.88 | 4.35 | 4.25 | 4.38 | 4.44 | 4.13 | 4.13 | 4.88 | 4.45 | 4.06 | 3.98 | 4.3834 | 4.04 | 4.375 | 4.61 | 4.223 | 4.39 | 3.09 | 4.29 | 4.0226 |  | 4.2498
| 1 | 5 | 4.33 | 4.4137 | 3.78 | 4.29 | 4.2 | 4.45 | 4.38 | 4.09 | 4.08 | 4.96 | 4.25 | 4.28 | 3.95 | 4.4045 | 4.11 | 4.368 | 4.61 | 3.988 | 4.36 | 2.65 | 4.38 | 3.9784 |  | 101
| 1 | 6 | 4.22 | 4.4758 | 3.83 | 4.26 | 4.2 | 4.44 | 4.31 | 4.01 | 3.9 | 4.99 | 4.53 | 4.35 | 3.91 | 4.5322 | 4.04 | 4.305 | 4.55 | 4.526 | 4.35 | 2.89 | 4.45 | 3.8618
| Mean |  | 4.2133 | 4.4711 | 3.8717 | 4.26 | 4.2367 | 4.4117 | 4.3883 | 4.075 | 4.0667 | 4.945 | 4.3917 | 4.2717 | 3.92 | 4.4222 | 4.065 | 4.327 | 4.6117 | 4.3175 | 4.3567 | 2.9533 | 4.35 | 3.9342
| Median |  | 4.2 | 4.4591 | 3.855 | 4.255 | 4.24 | 4.42 | 4.395 | 4.085 | 4.095 | 4.95 | 4.385 | 4.315 | 3.93 | 4.394 | 4.04 | 4.339 | 4.615 | 4.3435 | 4.355 | 3.015 | 4.335 | 3.9258
| Std Dev. |  | 0.069761 | 0.050953 | 0.094322 | 0.057271 | 0.03266 | 0.035449 | 0.043551 | 0.040866 | 0.083347 | 0.040866 | 0.094745 | 0.24112 | 0.053292 | 0.10123 | 0.057184 | 0.051722 | 0.033714 | 0.29848 | 0.025033 | 0.16597 | 0.057619 | 0.060404
| Rel.Std.Dev. |  | 1.656% | 1.14% | 2.436% | 1.344% | 0.7709% | 0.8035% | 0.9924% | 1.003% | 2.05% | 0.8264% | 2.157% | 5.645% | 1.359% | 2.289% | 1.407% | 1.195% | 0.7311% | 6.913% | 0.5746% | 5.62% | 1.325% | 1.535%
| PDM3 |  | -0.8574% | 5.208% | -8.897% | 0.2407% | -0.3083% | 3.81% | 3.261% | -4.112% | -4.309% | 16.36% | 3.339% | 0.5153% | -7.76% | 4.058% | -4.348% | 1.817% | 8.516% | 1.594% | 2.515% | -30.51% | 2.359% | -7.425%
| Z-Score (Absolute) |  | 0.45 | 0.97 | 2.33 | 0.19 | 0.32 | 0.65 | 0.52 | 1.21 | 1.26 | 3.58 | 0.53 | 0.13 | 2.06 | 0.7 | 1.27 | 0.18 | 1.75 | 0.13 | 0.34 | 7.39 | 0.31 | 1.99
| Analytical results for Sb in OREAS 931b (Certified Value 2.94 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 2
| 1 | 1 | < 5 | 2.7094 | 3.32 | 2.31 | 2.43 | 2.7 | 2.66 | 2.64 | 2.52 | 4.02 | < 5 | 3.71 | 2.82 | 3.29 | 3.18 | 1.8 | 4.21 | 3.18 | 3.0472 | 3.38 | < 1 | < 5 | 2.3247 | 1
| 1 | 2 | < 5 | 2.6111 | 3.36 | 2.41 | 2.35 | 2.68 | 2.72 | 2.65 | 2.55 | 3.92 | < 5 | 2.99 | 2.72 | 3.21 | 3.24 | 1.8 | 4.16 | 3.29 | 2.6875 | 3.26 | < 1 | < 5 | 2.4508 | 16
| 1 | 3 | < 5 | 2.8657 | 3.48 | 2.43 | 2.42 | 2.73 | 2.65 | 2.62 | 2.55 | 3.96 | < 5 | 3.32 | 2.73 | 3.2 | 2.93 | 1.7 | 4.25 | 3.33 | 2.9781 | 3.33 | < 1 | < 5 | 2.6363 | 16
| 1 | 4 | < 5 | 2.9636 | 3.35 | 2.56 | 2.51 | 2.67 | 2.63 | 2.65 | 2.61 | 3.68 | < 5 | 3.47 | 2.84 | 3.24 | 3 | 1.8 | 4.07 | 3.35 | 4.387 | 3.36 | < 1 | 6 | 2.6071 | 2.9442
| 1 | 5 | < 5 | 2.7901 | 3.2 | 2.31 | 2.48 | 2.73 | 2.71 | 2.64 | 2.58 | 3.6 | < 5 | 3.15 | 2.67 | 3.31 | 3.22 | 1.7 | 4.37 | 3.23 | 3.6872 | 3.29 | < 1 | < 5 | 2.5012 | 102
| 1 | 6 | < 5 | 2.825 | 3.34 | 2.3 | 2.51 | 2.69 | 2.75 | 2.61 | 2.51 | 3.58 | < 5 | 3.28 | 2.7 | 3.21 | 3.09 | 1.7 | 4.14 | 3.28 | 3.0653 | 3.46 | < 1 | < 5 | 2.5601
| Mean |  |  | 2.7942 | 3.3417 | 2.3867 | 2.45 | 2.7 | 2.6867 | 2.635 | 2.5533 | 3.7933 |  | 3.32 | 2.7467 | 3.2433 | 3.11 | 1.75 | 4.2 | 3.2767 | 3.3087 | 3.3467 |  | 6 | 2.5134
| Median |  |  | 2.8075 | 3.345 | 2.36 | 2.455 | 2.695 | 2.685 | 2.64 | 2.55 | 3.8 |  | 3.3 | 2.725 | 3.225 | 3.135 | 1.75 | 4.185 | 3.285 | 3.0562 | 3.345 |  | 6 | 2.5307
| Std Dev. |  |  | 0.12287 | 0.089536 | 0.10172 | 0.06229 | 0.025298 | 0.046762 | 0.016432 | 0.037238 | 0.19541 |  | 0.2506 | 0.06802 | 0.046332 | 0.12554 | 0.054772 | 0.10354 | 0.06314 | 0.62109 | 0.070899 |  |  | 0.11466
| Rel.Std.Dev. |  |  | 4.397% | 2.679% | 4.262% | 2.542% | 0.937% | 1.741% | 0.6236% | 1.458% | 5.152% |  | 7.548% | 2.476% | 1.429% | 4.037% | 3.13% | 2.465% | 1.927% | 18.77% | 2.118% |  |  | 4.562%
| PDM3 |  |  | -5.098% | 13.5% | -18.94% | -16.79% | -8.296% | -8.748% | -10.5% | -13.28% | 28.84% |  | 12.76% | -6.711% | 10.16% | 5.63% | -40.56% | 42.65% | 11.29% | 12.38% | 13.67% |  | 103.8% | -14.63%
| Z-Score (Absolute) |  | 0.46 | 0.09 | 1.11 | 0.67 | 0.56 | 0.09 | 0.11 | 0.21 | 0.36 | 1.96 | 0.46 | 1.07 | 0 | 0.93 | 0.68 | 1.87 | 2.72 | 0.99 | 1.05 | 1.12 | 4.21 | 0.63 | 0.44
Results from laboratories 17 and 23 were removed due to their high Z score reading.
| Analytical results for Sc in OREAS 931b (Certified Value 1.47 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  |  | 2
| 1 | 1 | < 5 | 1.4335 | 1.44 | 1 | 1.6 | 1.3 | 1.6 | 1.3 | 1.3 | 1.4 | 1.4 | 1.6 | 1.3 | 1.8 | 1.2 | 1.5 | 1.6 | 1.3598 | 1.6 | 1 | 1.3 |  |  | 1
| 1 | 2 | < 5 | 1.373 | 1.44 | 1 | 1.6 | 1.3 | 1.6 | 1.3 | 1.3 | 1.4 | 1.6 | 1.8 | 1.3 | 1.9 | 1.3 | 1.4 | 1.6 | 1.9247 | 1.5 | 1 | 1.3 |  |  | 0
| 1 | 3 | < 5 | 1.4403 | 1.46 | 1 | 1.5 | 1.4 | 1.5 | 1.3 | 1.4 | 1.4 | 1.6 | 1.8 | 1.3 | 1.9 | 1.2 | 1.5 | 1.6 | 1.8005 | 1.7 | 1 | 1.3 |  |  | 16
| 1 | 4 | < 5 | 1.4234 | 1.48 | 1 | 1.6 | 1.3 | 1.5 | 1.3 | 1.3 | 1.4 | 1.6 | 1.6 | 1.3 | 1.8 | 1.2 | 1.4 | 1.6 | 1.4719 | 1.7 | 1 | 1.3 |  |  | 1.4727
| 1 | 5 | < 5 | 1.2749 | 1.48 | 1 | 1.5 | 1.4 | 1.6 | 1.3 | 1.3 | 1.3 | 1.5 | 1.7 | 1.4 | 1.7 | 1.2 | 1.5 | 1.5 | 1.655 | 1.6 | 1 | 1.3 |  |  | 103
| 1 | 6 | < 5 | 1.2571 | 1.5 | 1 | 1.5 | 1.4 | 1.5 | 1.3 | 1.4 | 1.4 | 1.6 | 1.7 | 1.4 | 1.8 | 1.2 | 1.4 | 1.5 | 2.0991 | 1.7 | 1 | 1.3
| Mean |  |  | 1.367 | 1.4667 | 1 | 1.55 | 1.35 | 1.55 | 1.3 | 1.3333 | 1.3833 | 1.55 | 1.7 | 1.3333 | 1.8167 | 1.2167 | 1.45 | 1.5667 | 1.7185 | 1.6333 | 1 | 1.3
| Median |  |  | 1.3982 | 1.47 | 1 | 1.55 | 1.35 | 1.55 | 1.3 | 1.3 | 1.4 | 1.6 | 1.7 | 1.3 | 1.8 | 1.2 | 1.45 | 1.6 | 1.7278 | 1.65 | 1 | 1.3
| Std Dev. |  |  | 0.081944 | 0.024221 | 0 | 0.054772 | 0.054772 | 0.054772 | 0 | 0.05164 | 0.040825 | 0.083666 | 0.089443 | 0.05164 | 0.075277 | 0.040825 | 0.054772 | 0.05164 | 0.27843 | 0.08165 | 0 | 0
| Rel.Std.Dev. |  |  | 5.994% | 1.651% | 0% | 3.534% | 4.057% | 3.534% | 0% | 3.873% | 2.951% | 5.398% | 5.261% | 3.873% | 4.144% | 3.355% | 3.777% | 3.296% | 16.2% | 4.999% | 0% | 0%
| PDM3 |  |  | -7.177% | -0.4128% | -32.1% | 5.246% | -8.334% | 5.246% | -11.73% | -9.466% | -6.071% | 5.246% | 15.43% | -9.466% | 23.35% | -17.39% | -1.544% | 6.377% | 16.69% | 10.9% | -32.1% | -11.73%
| Z-Score (Absolute) |  | 5.23 | 0.5 | 0 | NA | 0.42 | 0.59 | 0.42 | 0.84 | 0.67 | 0.42 | 0.42 | 1.18 | 0.67 | 1.77 | 1.26 | 0.08 | 0.51 | 1.27 | 0.84 | NA | 0.84
Results from laboratories 4 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 931b (Certified Value 46.3 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 24 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.5g |  |  | 2
| 1 | 1 | 46 | 47.804 | 39.22 | 43 | 48.3 | 47.6 | 47 | 44.2 | 48.5 | 71 | 34.486 | 42 | 42.4 | 46.9 | 50 | 43.3 | 48.5 | 48 | 50 | 45 | 45.1106 |  |  | 1
| 1 | 2 | 46 | 47.8753 | 40.2 | 43 | 48.2 | 47.3 | 46.6 | 44.1 | 49.1 | 73 | 34.292 | 42 | 45.8 | 46.2 | 49 | 44.2 | 46.9 | 50.2 | 50 | 42 | 46.3578 |  |  | 0
| 1 | 3 | 46 | 48.9974 | 40.3 | 44 | 46.9 | 48.9 | 46.8 | 42.7 | 48.9 | 73 | 35.252 | 44 | 49.6 | 48.1 | 49 | 41.8 | 47.8 | 50.5 | 52 | 40 | 41.8831 |  |  | 16
| 1 | 4 | 46 | 49.1549 | 40.4 | 44 | 47.9 | 48.4 | 47.1 | 41.7 | 50.6 | 68 | 37.298 | 44 | 42.8 | 48.1 | 49 | 42.6 | 45.3 | 50.4 | 51 | 43 | 48.6012 |  |  | 46.3087
| 1 | 5 | 46 | 48.8636 | 40.89 | 43 | 47.5 | 47.5 | 46.8 | 42.5 | 50 | 73 | 35.636 | 42 | 46.3 | 48.2 | 48 | 42.9 | 48 | 49.4 | 52 | 42 | 48.3178 |  |  | 104
| 1 | 6 | 46 | 48.5748 | 42.26 | 46 | 47.8 | 50 | 46.6 | 42.2 | 49.2 | 71 | 38.352 | 45 | 45 | 48.6 | 49 | 43.6 | 47.2 | 49 | 52 | 44 | 49.0963
| Mean |  | 46 | 48.545 | 40.545 | 43.8333 | 47.7667 | 48.2833 | 46.8167 | 42.9 | 49.3833 | 71.5 | 35.886 | 43.1667 | 45.3167 | 47.6833 | 49 | 43.0667 | 47.2833 | 49.5833 | 51.1667 | 42.6667 | 46.5611
| Median |  | 46 | 48.7192 | 40.35 | 43.5 | 47.85 | 48 | 46.8 | 42.6 | 49.15 | 72 | 35.444 | 43 | 45.4 | 48.1 | 49 | 43.1 | 47.5 | 49.8 | 51.5 | 42.5 | 47.3378
| Std Dev. |  | 0 | 0.57906 | 1.0016 | 1.169 | 0.51251 | 1.0381 | 0.20412 | 1.0257 | 0.77309 | 1.9748 | 1.6149 | 1.3292 | 2.6263 | 0.92394 | 0.63246 | 0.83347 | 1.1268 | 0.97656 | 0.98319 | 1.7512 | 2.7461
| Rel.Std.Dev. |  | 0% | 1.193% | 2.47% | 2.667% | 1.073% | 2.15% | 0.436% | 2.391% | 1.565% | 2.762% | 4.5% | 3.079% | 5.796% | 1.938% | 1.291% | 1.935% | 2.383% | 1.97% | 1.922% | 4.104% | 5.898%
| PDM3 |  | -0.6665% | 4.829% | -12.45% | -5.345% | 3.148% | 4.264% | 1.097% | -7.361% | 6.64% | 54.4% | -22.51% | -6.785% | -2.142% | 2.969% | 5.812% | -7.001% | 2.105% | 7.071% | 10.49% | -7.865% | 0.5452%
| Z-Score (Absolute) |  | 0.21 | 0.45 | 1.65 | 0.78 | 0.25 | 0.39 | 0 | 1.03 | 0.67 | 6.48 | 2.87 | 0.96 | 0.39 | 0.23 | 0.57 | 0.99 | 0.12 | 0.73 | 1.14 | 1.09 | 0.07
| Analytical results for Sm in OREAS 931b (Indicative Value 3.83 ppm) | [24 empty] | Indicative
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  |  | 1
| No | No | 01: 03; 02: 15; 03: 17 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS |  |  | 2
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g |  |  | 2
| 1 | 1 | 01: 3.141; 02: 5.01; 03: 3.5 |  |  | 1
| 1 | 2 | 01: 3.25; 02: 4.99; 03: 3.5 |  |  | 17
| 1 | 3 | 01: 3.173; 02: 5.19; 03: 3.7 |  |  | 16
| 1 | 4 | 01: 3.199; 02: 4.98; 03: 3.1 |  |  | 3.8337
| 1 | 5 | 01: 3.16; 02: 4.7; 03: 3.1 |  |  | 11
| 1 | 6 | 01: 3.243; 02: 4.97; 03: 3.1
| Mean |  | 01: 3.1943; 02: 4.9733; 03: 3.3333
| Median |  | 01: 3.186; 02: 4.985; 03: 3.3
| Std Dev. |  | 01: 0.044648; 02: 0.15706; 03: 0.26583
| Rel.Std.Dev. |  | 01: 1.398%; 02: 3.158%; 03: 7.975%
| PDM3 |  | 01: -16.68%; 02: 29.73%; 03: -13.05%
| Z-Score (Absolute) |  | 01: 0.67; 02: 7.96; 03: 0
| Analytical results for Sn in OREAS 931b (Certified Value 34.6 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 23 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  |  |  | 1
| 1 | 1 | 36 | 30.5429 | 34.55 | 30.7 | 34.1 | 31.4 | 36.4 | 34.2 | 34.5 | 42.1 | 29.24 | 33.9 | 34.8 | 35.12 | 32.4 | 39.8 | 26.1711 | 40.2 | 25 | 36 |  |  |  | 1
| 1 | 2 | 36 | 30.8555 | 34.33 | 30.5 | 33.2 | 31.3 | 36.3 | 33.9 | 34.5 | 41.8 | 25.95 | 33.5 | 34.8 | 34.76 | 31.9 | 41.7 | 17.5587 | 38 | 25 | 35 |  |  |  | 8
| 1 | 3 | 36 | 31.2301 | 34.46 | 30.9 | 32.6 | 31.4 | 35.7 | 32.8 | 35.4 | 42.7 | 28.39 | 34.7 | 34.9 | 35.64 | 32.4 | 41.8 | 31.5661 | 41.6 | 24 | 36 |  |  |  | 16
| 1 | 4 | 36 | 31.1037 | 34.32 | 29.8 | 34.2 | 30.7 | 35.7 | 33.2 | 35.1 | 41.3 | 29.96 | 36.5 | 34.9 | 34.82 | 31.6 | 41.3 | 24.7481 | 40.7 | 25 | 38 |  |  |  | 34.5782
| 1 | 5 | 36 | 30.8258 | 34.11 | 31 | 33.1 | 31.8 | 35.6 | 33.1 | 35.7 | 41.4 | 29.27 | 33.1 | 35.8 | 33.81 | 32.7 | 39.7 | 24.9226 | 40 | 25 | 37 |  |  |  | 105
| 1 | 6 | 36 | 30.5412 | 34.48 | 31.3 | 33.5 | 31.9 | 33.8 | 33 | 35.2 | 41.2 | 27.7 | 36.5 | 34.5 | 33.82 | 32 | 40.5 | 25.2986 | 40.7 | 25 | 34
| Mean |  | 36 | 30.8499 | 34.375 | 30.7 | 33.45 | 31.4167 | 35.5833 | 33.3667 | 35.0667 | 41.75 | 28.4183 | 34.7 | 34.95 | 34.6617 | 32.1667 | 40.8 | 25.0442 | 40.2 | 24.8333 | 36
| Median |  | 36 | 30.8406 | 34.395 | 30.8 | 33.35 | 31.4 | 35.7 | 33.15 | 35.15 | 41.6 | 28.815 | 34.3 | 34.85 | 34.79 | 32.2 | 40.9 | 25.1106 | 40.45 | 25 | 36
| Std Dev. |  | 0 | 0.28252 | 0.15758 | 0.51769 | 0.61563 | 0.42622 | 0.93684 | 0.55377 | 0.48442 | 0.57533 | 1.4413 | 1.4913 | 0.44159 | 0.72604 | 0.40332 | 0.93381 | 4.4722 | 1.2116 | 0.40825 | 1.4142
| Rel.Std.Dev. |  | 0% | 0.9158% | 0.4584% | 1.686% | 1.84% | 1.357% | 2.633% | 1.66% | 1.381% | 1.378% | 5.072% | 4.298% | 1.263% | 2.095% | 1.254% | 2.289% | 17.86% | 3.014% | 1.644% | 3.928%
| PDM3 |  | 4.112% | -10.78% | -0.5877% | -11.22% | -3.263% | -9.143% | 2.907% | -3.504% | 1.413% | 20.74% | -17.81% | 0.3522% | 1.075% | 0.2413% | -6.974% | 17.99% | -27.57% | 16.26% | -28.18% | 4.112%
| Z-Score (Absolute) |  | NA | 1.08 | 0 | 1.12 | 0.28 | 0.9 | 0.37 | 0.31 | 0.21 | 2.25 | 1.82 | 0.1 | 0.18 | 0.09 | 0.67 | 1.96 | 2.85 | 1.78 | NA | NA
Results from laboratories 1, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 931b (Certified Value 13.5 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 1
| 1 | 1 | 14 | 17.8993 | 12.27 | 17.5 | 14.4 | 12.1 | 13.2 | 10.9 | 11.3 | < 0.5 | 13.26 | 9.4 | 14.5 | 11.6 | 21 | 11 | 18.3 | 11.4 | 16.2258 | 11.3 | 8 | 13.2 |  | 1
| 1 | 2 | 28 | 16.7695 | 12.21 | 17.5 | 14.2 | 12.2 | 13.6 | 11 | 12.5 | < 0.5 | 13.34 | 9.8 | 17 | 11.5 | 21 | 11 | 17.8 | 12.2 | 15.1467 | 11.6 | 9 | 12.8 |  | 0
| 1 | 3 | 28 | 18.1281 | 12.14 | 17.4 | 13.9 | 12.4 | 12.8 | 11 | 12.2 | < 0.5 | 13.15 | 9.7 | 16.8 | 11.4 | 20 | 10 | 18.2 | 11.8 | 15.5628 | 11.7 | 8 | 12 |  | 16
| 1 | 4 | 28 | 16.0361 | 12.15 | 17.7 | 14.2 | 12.2 | 12.7 | 11.4 | 12 | < 0.5 | 13.67 | 9.7 | 15.4 | 12.1 | 21 | 10 | 17.5 | 10.9 | 17.4648 | 12.8 | 8 | 13 |  | 13.4925
| 1 | 5 | 28 | 16.7294 | 11.87 | 17.6 | 14.2 | 12.4 | 13.4 | 11.2 | 12 | < 0.5 | 13.05 | 9.2 | 15.9 | 11.9 | 21 | 11 | 17.6 | 10.5 | 13.8759 | 12.4 | 9 | 12.7 |  | 106
| 1 | 6 | 27 | 16.0383 | 12.06 | 17.5 | 14.3 | 12.4 | 14 | 10.8 | 11.6 | < 0.5 | 13.16 | 10.3 | 16.5 | 11.2 | 21 | 10 | 18.8 | 11.5 | 13.7091 | 12.5 | 8 | 12.7
| Mean |  | 25.5 | 16.9334 | 12.1167 | 17.5333 | 14.2 | 12.2833 | 13.2833 | 11.05 | 11.9333 |  | 13.2717 | 9.6833 | 16.0167 | 11.6167 | 20.8333 | 10.5 | 18.0333 | 11.3833 | 15.3308 | 12.05 | 8.3333 | 12.7333
| Median |  | 28 | 16.7494 | 12.145 | 17.5 | 14.2 | 12.3 | 13.3 | 11 | 12 |  | 13.21 | 9.7 | 16.2 | 11.55 | 21 | 10.5 | 18 | 11.45 | 15.3547 | 12.05 | 8 | 12.75
| Std Dev. |  | 5.648 | 0.89835 | 0.13995 | 0.10328 | 0.16733 | 0.13292 | 0.4916 | 0.21679 | 0.4274 |  | 0.21904 | 0.37639 | 0.94956 | 0.33116 | 0.40825 | 0.54772 | 0.49261 | 0.61128 | 1.4276 | 0.59582 | 0.5164 | 0.40825
| Rel.Std.Dev. |  | 22.15% | 5.305% | 1.155% | 0.589% | 1.178% | 1.082% | 3.701% | 1.962% | 3.582% |  | 1.65% | 3.887% | 5.929% | 2.851% | 1.96% | 5.216% | 2.732% | 5.37% | 9.312% | 4.945% | 6.197% | 3.206%
| PDM3 |  | 88.99% | 25.5% | -10.2% | 29.95% | 5.243% | -8.962% | -1.55% | -18.1% | -11.56% |  | -1.637% | -28.23% | 18.71% | -13.9% | 54.41% | -22.18% | 33.65% | -15.63% | 13.62% | -10.69% | -38.24% | -5.627%
| Z-Score (Absolute) |  | NA | 1.93 | 0.28 | 2.21 | 0.67 | 0.21 | 0.25 | 0.77 | 0.37 | 5.74 | 0.25 | 1.4 | 1.51 | 0.51 | 3.72 | NA | 2.44 | 0.62 | 1.19 | 0.31 | NA | 0
Results from laboratories 1, 16 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 931b (Certified Value < 0.01 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | [6 empty] | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [6 empty] | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | [6 empty] | 3
| 1 | 1 | < 5 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.1 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.1 | < 0.05 | [6 empty] | 1
| 1 | 2 | < 5 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.1 | < 0.05 | [6 empty] | 10
| 1 | 3 | < 5 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.1 | < 0.05 | [6 empty] | 16
| 1 | 4 | < 5 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.1 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | < 0.1 | < 0.05 | [6 empty] | < 0.01
| 1 | 5 | < 5 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.09 | 0.06 | < 0.01 | < 0.05 | 0.01 | < 0.1 | < 0.05 | [6 empty] | 107
| 1 | 6 | < 5 | < 0.05 | < 0.01 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | 0.09 | 0.04 | < 0.01 | < 0.05 | < 0.01 | < 0.1 | < 0.05
| Mean |  | 12: 0.096667; 14: 0.036667; 18: 0.01
| Median |  | 12: 0.1; 14: 0.04; 18: 0.01
| Std Dev. |  | 12: 0.005164; 14: 0.025166
| Rel.Std.Dev. |  | 12: 5.342%; 14: 68.63%
PDM3
| Z-Score (Absolute) |  | 105.58 | 0.18 | 0.67 | 0.18 | 0.67 | 0.67 | 0.67 | 0.67 | 0.67 | 0.18 | 3.23 | 0 | 0.67 | 0.18 | 0.64 | 1.24 | 0.18
| Analytical results for Tb in OREAS 931b (Certified Value 0.36 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [6 empty] | 1
| No | No | 01: 03; 02: 10; 03: 12; 04: 15; 05: 17; 06: 20 | [6 empty] | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS | [6 empty] | 2
|  |  | 01: 0.5g; 02: 0.25g; 03: 0.25g; 04: 01g; 05: 0.5g; 06: 0.25g | [6 empty] | 2
| 1 | 1 | 01: 0.289; 02: 0.42; 03: 0.33; 04: 0.43; 05: 0.3; 06: 0.33 | [6 empty] | 1
| 1 | 2 | 01: 0.294; 02: 0.38; 03: 0.33; 04: 0.43; 05: 0.3; 06: 0.33 | [6 empty] | 0
| 1 | 3 | 01: 0.295; 02: 0.37; 03: 0.35; 04: 0.44; 05: 0.3; 06: 0.32 | [6 empty] | 16
| 1 | 4 | 01: 0.302; 02: 0.39; 03: 0.36; 04: 0.43; 05: 0.3; 06: 0.36 | [6 empty] | 0.35897
| 1 | 5 | 01: 0.288; 02: 0.4; 03: 0.34; 04: 0.4; 05: 0.3; 06: 0.33 | [6 empty] | 108
| 1 | 6 | 01: 0.301; 02: 0.4; 03: 0.37; 04: 0.42; 05: 0.3; 06: 0.34
| Mean |  | 01: 0.29483; 02: 0.39333; 03: 0.34667; 04: 0.425; 05: 0.3; 06: 0.335
| Median |  | 01: 0.2945; 02: 0.395; 03: 0.345; 04: 0.43; 05: 0.3; 06: 0.33
| Std Dev. |  | 01: 0.0058452; 02: 0.017512; 03: 0.01633; 04: 0.013784; 05: 0; 06: 0.013784
| Rel.Std.Dev. |  | 01: 1.983%; 02: 4.452%; 03: 4.711%; 04: 3.243%; 05: 0%; 06: 4.115%
| PDM3 |  | 01: -17.87%; 02: 9.574%; 03: -3.427%; 04: 18.4%; 05: -16.43%; 06: -6.677%
| Z-Score (Absolute) |  | 01: 0.75; 02: 0.67; 03: 0; 04: 1.13; 05: NA; 06: 0.17
Results from laboratory 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 931b (Certified Value 0.16 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 11 | 0.085764 | 0.16 | 0.13 | 0.17 | 0.15 | 0.14 | 0.16 | 0.15 | 0.28 | < 10 | 0.38 | 0.14 | 0.16 | 0.19 | < 0.2 | 0.19 | 0.17 | < 0.1 | 0.19 | 0.162 |  |  | 1
| 1 | 2 | 12 | 0.078914 | 0.17 | 0.14 | 0.17 | 0.17 | 0.18 | 0.17 | 0.16 | 0.29 | < 10 | 0.36 | 0.15 | 0.16 | 0.18 | < 0.2 | 0.15 | 0.17 | < 0.1 | 0.19 | 0.14919 |  |  | 21
| 1 | 3 | 12 | 0.09007 | 0.16 | 0.16 | 0.14 | 0.16 | 0.16 | 0.16 | 0.17 | 0.29 | < 10 | 0.4 | 0.16 | 0.17 | 0.19 | < 0.2 | 0.18 | 0.16 | < 0.1 | 0.18 | 0.18388 |  |  | 16
| 1 | 4 | 11 | 0.090812 | 0.16 | 0.17 | 0.16 | 0.16 | 0.16 | 0.14 | 0.16 | 0.25 | < 10 | 0.36 | 0.15 | 0.16 | 0.16 | < 0.2 | 0.15 | 0.17 | < 0.1 | 0.16 | 0.17794 |  |  | 0.16403
| 1 | 5 | 11 | 0.091209 | 0.16 | 0.14 | 0.16 | 0.16 | 0.17 | 0.15 | 0.18 | 0.32 | < 10 | 0.34 | 0.15 | 0.3 | 0.18 | < 0.2 | 0.17 | 0.16 | < 0.1 | 0.18 | 0.16497 |  |  | 109
| 1 | 6 | 12 | 0.076645 | 0.16 | 0.16 | 0.15 | 0.18 | 0.16 | 0.16 | 0.18 | 0.28 | < 10 | 0.35 | 0.16 | 0.21 | 0.19 | < 0.2 | 0.18 | 0.16 | < 0.1 | 0.18 | 0.16588
| Mean |  | 11.5 | 0.085569 | 0.16167 | 0.15 | 0.15833 | 0.16333 | 0.16167 | 0.15667 | 0.16667 | 0.285 |  | 0.365 | 0.15167 | 0.19333 | 0.18167 |  | 0.17 | 0.165 |  | 0.18 | 0.16731
| Median |  | 11.5 | 0.087917 | 0.16 | 0.15 | 0.16 | 0.16 | 0.16 | 0.16 | 0.165 | 0.285 |  | 0.36 | 0.15 | 0.165 | 0.185 |  | 0.175 | 0.165 |  | 0.18 | 0.16542
| Std Dev. |  | 0.54772 | 0.0063799 | 0.0040825 | 0.015492 | 0.01169 | 0.010328 | 0.013292 | 0.010328 | 0.012111 | 0.022583 |  | 0.021679 | 0.0075277 | 0.055737 | 0.01169 |  | 0.016733 | 0.0054772 |  | 0.010954 | 0.012262
| Rel.Std.Dev. |  | 4.763% | 7.456% | 2.525% | 10.33% | 7.383% | 6.323% | 8.222% | 6.592% | 7.266% | 7.924% |  | 5.94% | 4.963% | 28.83% | 6.435% |  | 9.843% | 3.32% |  | 6.086% | 7.329%
| PDM3 |  | 6911% | -47.83% | -1.443% | -8.556% | -3.475% | -0.4272% | -1.443% | -4.491% | 1.605% | 73.74% |  | 122.5% | -7.54% | 17.86% | 10.75% |  | 3.637% | 0.5888% |  | 9.733% | 1.997%
| Z-Score (Absolute) |  | NA | 3.97 | 0.13 | 0.72 | 0.3 | 0.04 | 0.13 | 0.38 | 0.13 | 6.11 | 244.56 | 10.16 | 0.63 | 1.48 | 0.89 | 3.25 | 0.3 | 0.04 | 5.77 | 0.8 | 0.16
Results from laboratory 1 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 931b (Certified Value 7.68 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | [5 empty] | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | [5 empty] | 3
| 1 | 1 | 7.7085 | 7.209 | 8.2 | 8.3 | 7.4 | 7.9 | 7.3 | 6.7 | 7.6 | 7 | 7.7 | 7.6 | 9.21 | 7 | 7.9 | 7.9 | 7.3341 | 7.7 | [5 empty] | 1
| 1 | 2 | 7.6273 | 7.312 | 8.05 | 7.8 | 7.4 | 8 | 7 | 6.7 | 7.5 | 6.88 | 8.8 | 7.6 | 9 | 7.4 | 7.9 | 7.8 | 7.4971 | 7.5 | [5 empty] | 11
| 1 | 3 | 7.7389 | 7.356 | 8.39 | 7.8 | 7.5 | 7.6 | 6.9 | 6.8 | 7.5 | 7.54 | 9.1 | 7.7 | 9.29 | 6.8 | 8.1 | 8.2 | 7.1726 | 7.6 | [5 empty] | 16
| 1 | 4 | 7.7679 | 7.483 | 8.43 | 8.2 | 7.1 | 7.9 | 7 | 6.9 | 7.4 | 7.6 | 8.4 | 7.8 | 8.96 | 6.8 | 7.8 | 8 | 7.4561 | 8 | [5 empty] | 7.6785
| 1 | 5 | 7.8176 | 7.276 | 8.33 | 8 | 7.6 | 8 | 7.1 | 6.8 | 7.5 | 7.02 | 8.5 | 7.9 | 8.53 | 7.1 | 8.3 | 7.6 | 7.156 | 7.6 | [5 empty] | 110
| 1 | 6 | 7.603 | 7.403 | 8.37 | 8.1 | 7.7 | 7.4 | 7 | 6.7 | 7.6 | 7.52 | 8.6 | 7.7 | 8.76 | 7.1 | 7.9 | 8.1 | 6.8632 | 7.9
| Mean |  | 7.7105 | 7.3398 | 8.295 | 8.0333 | 7.45 | 7.8 | 7.05 | 6.7667 | 7.5167 | 7.26 | 8.5167 | 7.7167 | 8.9583 | 7.0333 | 7.9833 | 7.9333 | 7.2465 | 7.7167
| Median |  | 7.7237 | 7.334 | 8.35 | 8.05 | 7.45 | 7.9 | 7 | 6.75 | 7.5 | 7.27 | 8.55 | 7.7 | 8.98 | 7.05 | 7.9 | 7.95 | 7.2533 | 7.65
| Std Dev. |  | 0.082511 | 0.096659 | 0.14363 | 0.20656 | 0.20736 | 0.24495 | 0.13784 | 0.08165 | 0.075277 | 0.32594 | 0.47081 | 0.1169 | 0.28209 | 0.22509 | 0.18348 | 0.21602 | 0.23451 | 0.19408
| Rel.Std.Dev. |  | 1.07% | 1.317% | 1.732% | 2.571% | 2.783% | 3.14% | 1.955% | 1.207% | 1.001% | 4.49% | 5.528% | 1.515% | 3.149% | 3.2% | 2.298% | 2.723% | 3.236% | 2.515%
| PDM3 |  | 0.4168% | -4.411% | 8.028% | 4.621% | -2.976% | 1.582% | -8.186% | -11.88% | -2.108% | -5.451% | 10.92% | 0.4967% | 16.67% | -8.403% | 3.97% | 3.318% | -5.626% | 0.4967%
| Z-Score (Absolute) |  | 0.01 | 0.73 | 1.13 | 0.62 | 0.51 | 0.17 | 1.29 | 1.84 | 0.38 | 0.88 | 1.56 | 0.01 | 2.42 | 1.32 | 0.52 | 0.43 | 0.91 | 0.01
| Analytical results for Ti in OREAS 931b (Certified Value 0.007 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  |  | wt.%
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES |  |  | 3
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  |  | 3
| 1 | 1 | 0.01 | 0.0097624 | 0.0044 | 0.0081 | 0.006 | 0.006 | 0.006 | 0.005 | 0.006 | < 0.01 | 0.005 | 0.009 | 0.005 | 0.0086 | 0.005 | < 0.01 | < 0.005 | 0.008369 | < 0.01 | < 0.006 | 0.007 |  |  | 1
| 1 | 2 | 0.01 | 0.0096075 | 0.0044 | 0.0078 | 0.006 | 0.006 | 0.007 | 0.005 | 0.006 | < 0.01 | 0.005 | 0.01 | 0.005 | 0.0083 | 0.004 | < 0.01 | < 0.005 | 0.005081 | < 0.01 | < 0.006 | 0.007 |  |  | 0
| 1 | 3 | 0.01 | 0.0097805 | 0.0043 | 0.008 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | < 0.01 | 0.005 | 0.009 | 0.005 | 0.0074 | 0.004 | < 0.01 | < 0.005 | 0.008267 | < 0.01 | < 0.006 | 0.007 |  |  | 16
| 1 | 4 | 0.01 | 0.0084559 | 0.0043 | 0.0082 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | < 0.01 | 0.005 | 0.009 | 0.006 | 0.0088 | 0.004 | < 0.01 | < 0.005 | 0.008697 | < 0.01 | < 0.006 | 0.007 |  |  | 0.006777
| 1 | 5 | 0.01 | 0.0087427 | 0.0044 | 0.0085 | 0.006 | 0.006 | 0.006 | 0.006 | 0.006 | < 0.01 | 0.005 | 0.009 | 0.006 | 0.0086 | 0.004 | < 0.01 | < 0.005 | 0.007207 | < 0.01 | < 0.006 | 0.007 |  |  | 111
| 1 | 6 | 0.01 | 0.008357 | 0.0045 | 0.0078 | 0.006 | 0.006 | 0.007 | 0.005 | 0.005 | < 0.01 | 0.005 | 0.009 | 0.006 | 0.008 | 0.004 | < 0.01 | < 0.005 | 0.007866 | < 0.01 | < 0.006 | 0.007
| Mean |  | 0.01 | 0.0091177 | 0.0043833 | 0.0080667 | 0.006 | 0.006 | 0.0063333 | 0.0055 | 0.0058333 |  | 0.005 | 0.0091667 | 0.0055 | 0.0082833 | 0.0041667 |  |  | 0.0075812 |  |  | 0.007
| Median |  | 0.01 | 0.0091751 | 0.0044 | 0.00805 | 0.006 | 0.006 | 0.006 | 0.0055 | 0.006 |  | 0.005 | 0.009 | 0.0055 | 0.00845 | 0.004 |  |  | 0.0080665 |  |  | 0.007
| Std Dev. |  | 0 | 0.00067113 | 7.5277e-05 | 0.00026583 | 9.5015e-19 | 9.5015e-19 | 0.0005164 | 0.00054772 | 0.00040825 |  | 0 | 0.00040825 | 0.00054772 | 0.00051543 | 0.00040825 |  |  | 0.0013273 |  |  | 9.5015e-19
| Rel.Std.Dev. |  | 0% | 7.361% | 1.717% | 3.295% | 1.584e-14% | 1.584e-14% | 8.154% | 9.959% | 6.999% |  | 0% | 4.454% | 9.959% | 6.222% | 9.798% |  |  | 17.51% |  |  | 1.357e-14%
| PDM3 |  | 47.56% | 34.54% | -35.32% | 19.03% | -11.47% | -11.47% | -6.547% | -18.84% | -13.92% |  | -26.22% | 35.26% | -18.84% | 22.23% | -38.52% |  |  | 11.87% |  |  | 3.29%
| Z-Score (Absolute) |  | 2.41 | 1.9 | 0.84 | 1.29 | 0.1 | 0.1 | 0.29 | 0.19 | 0 | 0.48 | 0.48 | 1.93 | 0.19 | 1.42 | 0.96 | 0.48 | 1.93 | 1.01 | 0.48 | 1.64 | 0.67
| Analytical results for Tl in OREAS 931b (Certified Value 0.11 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS |  |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | < 10 | 0.11346 | 0.12 | 0.15 | 0.11 | 0.1 | 0.12 | 0.11 | 0.09 | 0.11 | 0.12 | 0.12 | < 0.01 | 0.18 | 0.1 | 0.11 | 0.1 | 0.14903 | 0.09 | 0.1035 |  |  |  | 1
| 1 | 2 | < 10 | 0.11666 | 0.12 | 0.15 | 0.11 | 0.1 | 0.12 | 0.11 | 0.1 | 0.11 | 0.1 | 0.14 | < 0.01 | 0.18 | 0.1 | 0.11 | 0.11 | 0.15435 | 0.09 | 0.1055 |  |  |  | 0
| 1 | 3 | < 10 | 0.12318 | 0.12 | 0.16 | 0.11 | 0.11 | 0.11 | 0.1 | 0.1 | 0.11 | 0.12 | 0.15 | < 0.01 | 0.19 | < 0.1 | 0.11 | 0.11 | 0.14457 | 0.09 | 0.1054 |  |  |  | 16
| 1 | 4 | < 10 | 0.12241 | 0.12 | 0.16 | 0.11 | 0.11 | 0.11 | 0.12 | 0.1 | 0.12 | 0.12 | 0.13 | < 0.01 | 0.16 | < 0.1 | 0.1 | 0.11 | 0.1449 | 0.1 | 0.1095 |  |  |  | 0.11103
| 1 | 5 | < 10 | 0.11556 | 0.12 | 0.16 | 0.12 | 0.11 | 0.11 | 0.11 | 0.1 | 0.12 | 0.11 | 0.13 | < 0.01 | 0.15 | < 0.1 | 0.12 | 0.1 | 0.13196 | 0.09 | 0.1099 |  |  |  | 112
| 1 | 6 | < 10 | 0.12027 | 0.11 | 0.15 | 0.11 | 0.11 | 0.11 | 0.1 | 0.1 | 0.12 | 0.12 | 0.14 | < 0.01 | 0.17 | < 0.1 | 0.11 | 0.11 | 0.12793 | 0.1 | 0.1092
| Mean |  |  | 0.11859 | 0.11833 | 0.155 | 0.11167 | 0.10667 | 0.11333 | 0.10833 | 0.098333 | 0.115 | 0.115 | 0.135 |  | 0.17167 | 0.1 | 0.11 | 0.10667 | 0.14212 | 0.093333 | 0.10717
| Median |  |  | 0.11847 | 0.12 | 0.155 | 0.11 | 0.11 | 0.11 | 0.11 | 0.1 | 0.115 | 0.12 | 0.135 |  | 0.175 | 0.1 | 0.11 | 0.11 | 0.14473 | 0.09 | 0.10735
| Std Dev. |  |  | 0.0039397 | 0.0040825 | 0.0054772 | 0.0040825 | 0.005164 | 0.005164 | 0.0075277 | 0.0040825 | 0.0054772 | 0.0083666 | 0.010488 |  | 0.01472 | 0 | 0.0063246 | 0.005164 | 0.010153 | 0.005164 | 0.0026979
| Rel.Std.Dev. |  |  | 3.322% | 3.45% | 3.534% | 3.656% | 4.841% | 4.556% | 6.949% | 4.152% | 4.763% | 7.275% | 7.769% |  | 8.575% | 0% | 5.75% | 4.841% | 7.144% | 5.533% | 2.517%
| PDM3 |  |  | 6.809% | 6.578% | 39.6% | 0.5732% | -3.93% | 2.074% | -2.429% | -11.44% | 3.575% | 3.575% | 21.59% |  | 54.61% | -9.934% | -0.9279% | -3.93% | 28% | -15.94% | -3.48%
| Z-Score (Absolute) |  | 494.27 | 0.53 | 0.51 | 4.21 | 0.17 | 0.67 | 0 | 0.51 | 1.52 | 0.17 | 0.17 | 2.19 | 10.96 | 5.9 | NA | 0.34 | 0.67 | 2.91 | 2.02 | 0.62
Results from laboratory 16 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 931b (Indicative Value 0.09 ppm) | [24 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab |  |  |  | 1
| No | No | 01: 03; 02: 15; 03: 17 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS |  |  |  | 3
|  |  | 01: 0.5g; 02: 01g; 03: 0.5g |  |  |  | 3
| 1 | 1 | 01: 0.077; 02: 0.11; 03: < 0.1 |  |  |  | 1
| 1 | 2 | 01: 0.077; 02: 0.1; 03: < 0.1 |  |  |  | 6
| 1 | 3 | 01: 0.082; 02: 0.11; 03: < 0.1 |  |  |  | 16
| 1 | 4 | 01: 0.077; 02: 0.1; 03: < 0.1 |  |  |  | 0.090083
| 1 | 5 | 01: 0.08; 02: 0.09; 03: < 0.1 |  |  |  | 12
| 1 | 6 | 01: 0.078; 02: 0.1; 03: < 0.1
| Mean |  | 01: 0.0785; 02: 0.10167
| Median |  | 01: 0.0775; 02: 0.1
| Std Dev. |  | 01: 0.0020736; 02: 0.0075277
| Rel.Std.Dev. |  | 01: 2.642%; 02: 7.404%
| PDM3 |  | 01: -12.86%; 02: 12.86%
| Z-Score (Absolute) |  | 01: 0; 02: 0.67; 03: 0.83
| Analytical results for U in OREAS 931b (Certified Value 2.08 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.5g |  |  |  | 3
| 1 | 1 | 2.0324 | 1.989 | 2.14 | 1.99 | 2.07 | 2.23 | 2.05 | 1.81 | 2.17 | 1.95 | 1.84 | 2.18 | 2.28 | 2 | 2 | 2.17 | 2.0481 | 2.1 | 8 | 1.9744 |  |  |  | 1
| 1 | 2 | 2.0683 | 2.011 | 2.16 | 1.96 | 2.05 | 2.17 | 2 | 1.8 | 2.16 | 1.95 | 2.08 | 2.17 | 2.27 | 2.1 | 2 | 2.21 | 2.1459 | 2.07 | 8 | 2.031 |  |  |  | 13
| 1 | 3 | 2.074 | 2.012 | 2.15 | 1.92 | 2.14 | 2.2 | 1.94 | 1.83 | 2.2 | 2.05 | 2.25 | 2.18 | 2.32 | 1.9 | 2.1 | 2.3 | 2.0464 | 2.14 | 8 | 2.0818 |  |  |  | 16
| 1 | 4 | 2.0293 | 2.041 | 2.25 | 2.02 | 2.04 | 2.22 | 1.97 | 1.92 | 2.11 | 2.15 | 2.04 | 2.25 | 2.26 | 1.9 | 2 | 2.3 | 2.0807 | 2.26 | 8 | 2.0488 |  |  |  | 2.0816
| 1 | 5 | 2.0391 | 2.022 | 2.21 | 1.97 | 2.11 | 2.18 | 1.98 | 1.89 | 2.16 | 2 | 2.06 | 2.18 | 2.15 | 2 | 2.2 | 2.11 | 2.0468 | 2.13 | 8 | 2.085 |  |  |  | 113
| 1 | 6 | 2.0438 | 2.034 | 2.24 | 1.96 | 2.15 | 2.05 | 1.96 | 1.81 | 2.2 | 2.1 | 2.1 | 2.15 | 2.17 | 2 | 2.1 | 2.24 | 1.9699 | 2.18 | 8 | 2.1008
| Mean |  | 2.0478 | 2.0182 | 2.1917 | 1.97 | 2.0933 | 2.175 | 1.9833 | 1.8433 | 2.1667 | 2.0333 | 2.0617 | 2.185 | 2.2417 | 1.9833 | 2.0667 | 2.2217 | 2.0563 | 2.1467 | 8 | 2.0536
| Median |  | 2.0414 | 2.017 | 2.185 | 1.965 | 2.09 | 2.19 | 1.975 | 1.82 | 2.165 | 2.025 | 2.07 | 2.18 | 2.265 | 2 | 2.05 | 2.225 | 2.0474 | 2.135 | 8 | 2.0653
| Std Dev. |  | 0.018835 | 0.018584 | 0.047924 | 0.033466 | 0.046762 | 0.065345 | 0.038297 | 0.049666 | 0.033267 | 0.08165 | 0.13182 | 0.033912 | 0.066758 | 0.075277 | 0.08165 | 0.074677 | 0.057157 | 0.066833 | 0 | 0.046509
| Rel.Std.Dev. |  | 0.9198% | 0.9208% | 2.187% | 1.699% | 2.234% | 3.004% | 1.931% | 2.694% | 1.535% | 4.016% | 6.394% | 1.552% | 2.978% | 3.795% | 3.951% | 3.361% | 2.78% | 3.113% | 0% | 2.265%
| PDM3 |  | -1.626% | -3.049% | 5.285% | -5.363% | 0.5615% | 4.485% | -4.723% | -11.45% | 4.084% | -2.321% | -0.9597% | 4.965% | 7.687% | -4.723% | -0.7195% | 6.727% | -1.217% | 3.124% | 284.3% | -1.346%
| Z-Score (Absolute) |  | 0.12 | 0.37 | 1.12 | 0.79 | 0.27 | 0.98 | 0.67 | 1.88 | 0.9 | 0.24 | 0 | 1.06 | 1.55 | 0.67 | 0.04 | 1.38 | 0.05 | 0.73 | NA | 0.07
Results from laboratory 22 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 931b (Certified Value 18.8 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  | 1
| 1 | 1 | 19 | 21.4546 | 18.6 | 22 | 18 | 17 | 18 | 16 | 17 | 19 | 18.42 | 10 | 19 | 15.5 | 24 | 18 | 22 | 18 | 21.2583 | 16 | 19 | 15 |  | 1
| 1 | 2 | 19 | 22.1788 | 18 | 21 | 18 | 17 | 18 | 16 | 17 | 18 | 17.88 | 11 | 23 | 15.2 | 24 | 18 | 21 | 18 | 21.2565 | 16 | 20 | 15 |  | 0
| 1 | 3 | 20 | 22.0114 | 18.4 | 22 | 18 | 18 | 18 | 17 | 17 | 19 | 18.34 | 11 | 23 | 15.4 | 22 | 18 | 21 | 18 | 21.5126 | 17 | 19 | 16 |  | 16
| 1 | 4 | 19 | 20.869 | 18.6 | 23 | 18 | 17 | 18 | 17 | 17 | 19 | 18.41 | 11 | 20 | 15.8 | 23 | 18 | 21 | 17 | 22.5163 | 17 | 19 | 16 |  | 18.7805
| 1 | 5 | 19 | 21.4208 | 18.3 | 23 | 18 | 17 | 18 | 17 | 16 | 19 | 18.33 | 10 | 21 | 15.8 | 24 | 18 | 22 | 17 | 20.5209 | 17 | 19 | 15 |  | 114
| 1 | 6 | 19 | 21.1531 | 18.4 | 22 | 18 | 18 | 18 | 16 | 16 | 19 | 18.32 | 11 | 22 | 14.9 | 23 | 18 | 22 | 17 | 20.5881 | 17 | 19 | 15
| Mean |  | 19.1667 | 21.5146 | 18.3833 | 22.1667 | 18 | 17.3333 | 18 | 16.5 | 16.6667 | 18.8333 | 18.2833 | 10.6667 | 21.3333 | 15.4333 | 23.3333 | 18 | 21.5 | 17.5 | 21.2754 | 16.6667 | 19.1667 | 15.3333
| Median |  | 19 | 21.4377 | 18.4 | 22 | 18 | 17 | 18 | 16.5 | 17 | 19 | 18.335 | 11 | 21.5 | 15.45 | 23.5 | 18 | 21.5 | 17.5 | 21.2574 | 17 | 19 | 15
| Std Dev. |  | 0.40825 | 0.49957 | 0.22286 | 0.75277 | 0 | 0.5164 | 0 | 0.54772 | 0.5164 | 0.40825 | 0.20206 | 0.5164 | 1.633 | 0.35024 | 0.8165 | 0 | 0.54772 | 0.54772 | 0.72641 | 0.5164 | 0.40825 | 0.5164
| Rel.Std.Dev. |  | 2.13% | 2.322% | 1.212% | 3.396% | 0% | 2.979% | 0% | 3.32% | 3.098% | 2.168% | 1.105% | 4.841% | 7.655% | 2.269% | 3.499% | 0% | 2.548% | 3.13% | 3.414% | 3.098% | 2.13% | 3.368%
| PDM3 |  | 2.056% | 14.56% | -2.115% | 18.03% | -4.156% | -7.706% | -4.156% | -12.14% | -11.26% | 0.2814% | -2.647% | -43.2% | 13.59% | -17.82% | 24.24% | -4.156% | 14.48% | -6.818% | 13.28% | -11.26% | 2.056% | -18.35%
| Z-Score (Absolute) |  | 0.47 | 1.54 | 0.11 | 1.84 | 0.06 | 0.37 | 0.06 | 0.75 | 0.67 | 0.32 | 0.06 | 3.42 | 1.46 | 1.24 | 2.37 | 0.06 | 1.54 | 0.29 | 1.43 | 0.67 | 0.47 | 1.28
| Analytical results for W in OREAS 931b (Certified Value 13.7 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES |  |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  |  | 2
| 1 | 1 | 16 | 14.2973 | 14.72 | 12.2 | 14.4 | 12.75 | 14.65 | 13.9 | 13.45 | 17.9 | 13.96 | 13.28 | 18.4 | 13.12 | 11.9 | 13.8 | 15.13 | 12.1835 | 13.7 | 18 | 16 |  |  | 1
| 1 | 2 | 16 | 14.4537 | 14.85 | 11.7 | 14.85 | 12.9 | 14.15 | 14.05 | 13.6 | 17.8 | 14.56 | 13.65 | 18.5 | 13.11 | 11.6 | 14 | 16.08 | 10.4912 | 13.3 | 19 | 19 |  |  | 15
| 1 | 3 | 16 | 15.1452 | 15.17 | 11.2 | 13.75 | 12.8 | 15.05 | 13.35 | 13.75 | 18.1 | 15.59 | 14.23 | 18.7 | 13.36 | 11.4 | 14.8 | 16.2 | 13.5074 | 13.4 | 18 | 15 |  |  | 16
| 1 | 4 | 17 | 14.9861 | 15.14 | 11.4 | 15.05 | 12.7 | 14.45 | 13.1 | 13.65 | 17.7 | 15.48 | 13.86 | 19.2 | 12.76 | 11 | 14.3 | 15.93 | 11.1492 | 13.8 | 18 | 17 |  |  | 13.6653
| 1 | 5 | 18 | 14.469 | 14.39 | 12.4 | 15.1 | 12.95 | 14.95 | 13.4 | 13.65 | 17.8 | 14.58 | 13.57 | 19.4 | 11.88 | 10.9 | 14.9 | 15.16 | 11.5387 | 13.1 | 19 | 18 |  |  | 115
| 1 | 6 | 17 | 14.5945 | 15.04 | 12.1 | 13.5 | 13.2 | 14 | 13.55 | 13.85 | 18.2 | 15.93 | 13.59 | 19.2 | 12.66 | 11.2 | 14.1 | 15.8 | 16.605 | 13.8 | 18 | 17
| Mean |  | 16.6667 | 14.6576 | 14.885 | 11.8333 | 14.4417 | 12.8833 | 14.5417 | 13.5583 | 13.6583 | 17.9167 | 15.0167 | 13.6967 | 18.9 | 12.815 | 11.3333 | 14.3167 | 15.7167 | 12.5792 | 13.5167 | 18.3333 | 17
| Median |  | 16.5 | 14.5318 | 14.945 | 11.9 | 14.625 | 12.85 | 14.55 | 13.475 | 13.65 | 17.85 | 15.03 | 13.62 | 18.95 | 12.935 | 11.3 | 14.2 | 15.865 | 11.8611 | 13.55 | 18 | 17
| Std Dev. |  | 0.8165 | 0.33363 | 0.29791 | 0.4761 | 0.68368 | 0.18074 | 0.42239 | 0.35695 | 0.13571 | 0.19408 | 0.7607 | 0.32073 | 0.41952 | 0.52504 | 0.37771 | 0.4446 | 0.46306 | 2.2232 | 0.29269 | 0.5164 | 1.4142
| Rel.Std.Dev. |  | 4.899% | 2.276% | 2.001% | 4.023% | 4.734% | 1.403% | 2.905% | 2.633% | 0.9936% | 1.083% | 5.066% | 2.342% | 2.22% | 4.097% | 3.333% | 3.105% | 2.946% | 17.67% | 2.165% | 2.817% | 8.319%
| PDM3 |  | 21.96% | 7.262% | 8.925% | -13.41% | 5.681% | -5.722% | 6.413% | -0.7828% | -0.05106% | 31.11% | 9.889% | 0.2295% | 38.31% | -6.222% | -17.06% | 4.766% | 15.01% | -7.948% | -1.088% | 34.16% | 24.4%
| Z-Score (Absolute) |  | NA | 0.46 | 0.63 | 1.55 | 0.31 | 0.8 | 0.38 | 0.32 | 0.25 | 2.79 | 0.72 | 0.22 | 3.49 | 0.85 | 1.91 | 0.22 | 1.22 | 1.02 | 0.35 | NA | NA
Results from laboratories 1, 22 and 23 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 931b (Certified Value 5.53 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g | 0.25g | 01g | 0.25g |  |  |  |  | 3
| 1 | 1 | 7 | 5.9471 | 5.23 | 6 | 6.11 | 5.26 | 5.66 | 4.92 | 5.08 | 6.02 | 5.24 | 4.71 | 4.86 | 6.89 | 5.93 | 5.3 | 5.51 | 7 | 5.3 |  |  |  |  | 1
| 1 | 2 | 7 | 6.0613 | 5.27 | 5.98 | 5.91 | 5.27 | 5.62 | 4.92 | 5.59 | 5.78 | 5.144 | 4.74 | 4.81 | 6.71 | 5.64 | 5.44 | 5.42 | 7 | 5.1 |  |  |  |  | 16
| 1 | 3 | 7 | 5.9802 | 5.32 | 6.15 | 5.73 | 5.45 | 5.55 | 4.81 | 5.3 | 5.79 | 5.275 | 5.09 | 4.86 | 7.01 | 5.65 | 5.43 | 5.94 | 6 | 5.3 |  |  |  |  | 16
| 1 | 4 | 7 | 5.8388 | 5.27 | 6.27 | 5.96 | 5.36 | 5.44 | 5.06 | 5.36 | 5.83 | 5.3 | 5.01 | 4.98 | 6.62 | 5.58 | 5.36 | 6.03 | 6 | 5.2 |  |  |  |  | 5.5287
| 1 | 5 | 7 | 5.7706 | 5.28 | 6.31 | 5.71 | 5.55 | 5.72 | 5.01 | 5.33 | 5.8 | 5.22 | 4.77 | 4.94 | 6.34 | 5.67 | 5.2 | 5.88 | 7 | 5.1 |  |  |  |  | 116
| 1 | 6 | 7 | 5.9244 | 5.37 | 6.13 | 5.97 | 5.5 | 5.85 | 4.83 | 5.2 | 5.91 | 5.245 | 5.21 | 4.77 | 6.61 | 5.76 | 5.33 | 5.99 | 6 | 5.1
| Mean |  | 7 | 5.9204 | 5.29 | 6.14 | 5.8983 | 5.3983 | 5.64 | 4.925 | 5.31 | 5.855 | 5.2373 | 4.9217 | 4.87 | 6.6967 | 5.705 | 5.3433 | 5.795 | 6.5 | 5.1833
| Median |  | 7 | 5.9357 | 5.275 | 6.14 | 5.935 | 5.405 | 5.64 | 4.92 | 5.315 | 5.815 | 5.2425 | 4.89 | 4.86 | 6.665 | 5.66 | 5.345 | 5.91 | 6.5 | 5.15
| Std Dev. |  | 0 | 0.10319 | 0.04858 | 0.13506 | 0.15342 | 0.1209 | 0.141 | 0.097724 | 0.17111 | 0.093541 | 0.053709 | 0.2098 | 0.078486 | 0.23509 | 0.1247 | 0.089144 | 0.26205 | 0.54772 | 0.098319
| Rel.Std.Dev. |  | 0% | 1.743% | 0.9183% | 2.2% | 2.601% | 2.24% | 2.5% | 1.984% | 3.222% | 1.598% | 1.026% | 4.263% | 1.612% | 3.511% | 2.186% | 1.668% | 4.522% | 8.427% | 1.897%
| PDM3 |  | 26.61% | 7.084% | -4.318% | 11.06% | 6.685% | -2.359% | 2.012% | -10.92% | -3.957% | 5.901% | -5.271% | -10.98% | -11.91% | 21.12% | 3.188% | -3.354% | 4.816% | 17.57% | -6.248%
| Z-Score (Absolute) |  | NA | 0.89 | 0.18 | 1.26 | 0.85 | 0 | 0.41 | 0.8 | 0.15 | 0.78 | 0.27 | 0.81 | 0.9 | 2.21 | 0.52 | 0.09 | 0.67 | NA | 0.37
Results from laboratories 1 and 22 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 931b (Certified Value 0.61 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  |  | 1
| No | No | 01: 03; 02: 10; 03: 12; 04: 15; 05: 17; 06: 20 |  |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS |  |  |  |  | 2
|  |  | 01: 0.5g; 02: 0.25g; 03: 0.25g; 04: 01g; 05: 0.5g; 06: 0.25g |  |  |  |  | 3
| 1 | 1 | 01: 0.588; 02: 0.6; 03: 0.6; 04: 0.67; 05: 0.6; 06: 0.6 |  |  |  |  | 1
| 1 | 2 | 01: 0.59; 02: 0.6; 03: 0.6; 04: 0.67; 05: 0.6; 06: 0.6 |  |  |  |  | 0
| 1 | 3 | 01: 0.572; 02: 0.6; 03: 0.6; 04: 0.67; 05: 0.6; 06: 0.6 |  |  |  |  | 16
| 1 | 4 | 01: 0.585; 02: 0.6; 03: 0.6; 04: 0.66; 05: 0.6; 06: 0.6 |  |  |  |  | 0.60808
| 1 | 5 | 01: 0.593; 02: 0.6; 03: 0.6; 04: 0.66; 05: 0.6; 06: 0.6 |  |  |  |  | 117
| 1 | 6 | 01: 0.573; 02: 0.6; 03: 0.6; 04: 0.66; 05: 0.6; 06: 0.6
| Mean |  | 01: 0.5835; 02: 0.6; 03: 0.6; 04: 0.665; 05: 0.6; 06: 0.6
| Median |  | 01: 0.5865; 02: 0.6; 03: 0.6; 04: 0.665; 05: 0.6; 06: 0.6
| Std Dev. |  | 01: 0.0089163; 02: 0; 03: 0; 04: 0.0054772; 05: 0; 06: 0
| Rel.Std.Dev. |  | 01: 1.528%; 02: 0%; 03: 0%; 04: 0.8236%; 05: 0%; 06: 0%
| PDM3 |  | 01: -4.043%; 02: -1.329%; 03: -1.329%; 04: 9.36%; 05: -1.329%; 06: -1.329%
Z-Score (Absolute) | bar: 01 to 06
| Analytical results for Zn in OREAS 931b (Certified Value 272 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | 0
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 0.5g | 01g | 0.5g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g | 0.5g | 0
| 1 | 1 | 272 | 281.0263 | 277.7 | 276 | 267 | 271 | 270 | 253 | 266 | 300 | 248.214 | 274 | 309 | 275 | 271 | 272 | 269 | 286 | 321.842 | 272 | 270 | 275 | 255.5071 | 1
| 1 | 2 | 274 | 277.5873 | 276.5 | 270 | 268 | 273 | 271 | 257 | 266 | 297 | 240.291 | 276 | 344.5 | 276 | 285 | 275 | 261 | 281 | 325.301 | 266 | 274 | 274 | 264.4367 | 0
| 1 | 3 | 272 | 282.3775 | 275 | 273 | 268 | 275 | 271 | 256 | 267 | 301 | 247.353 | 277 | 365.9 | 277 | 277 | 262 | 268 | 281 | 332.8307 | 267 | 264 | 275 | 237.9844 | 16
| 1 | 4 | 272 | 280.5078 | 276.8 | 278 | 271 | 275 | 274 | 259 | 268 | 294 | 248.259 | 284 | 314.8 | 280 | 275 | 262 | 258 | 278 | 333.1996 | 267 | 262 | 277 | 275.6268 | 271.94
| 1 | 5 | 272 | 278.6093 | 273.4 | 278 | 266 | 275 | 271 | 255 | 266 | 300 | 246.292 | 266 | 341.1 | 280 | 276 | 268 | 269 | 281 | 329.0337 | 267 | 275 | 277 | 276.9878 | 118
| 1 | 6 | 272 | 277.4842 | 274.7 | 277 | 263 | 277 | 272 | 249 | 253 | 300 | 249.027 | 288 | 343.6 | 271 | 284 | 270 | 265 | 279 | 322.9317 | 266 | 270 | 275 | 283.162
| Mean |  | 272.3333 | 279.5987 | 275.6833 | 275.3333 | 267.1667 | 274.3333 | 271.5 | 254.8333 | 264.3333 | 298.6667 | 246.5727 | 277.5 | 336.4833 | 276.5 | 278 | 268.1667 | 265 | 281 | 327.5231 | 267.5 | 269.1667 | 275.5 | 265.6175
| Median |  | 272 | 279.5585 | 275.75 | 276.5 | 267.5 | 275 | 271 | 255.5 | 266 | 300 | 247.7835 | 276.5 | 342.35 | 276.5 | 276.5 | 269 | 266.5 | 281 | 327.1673 | 267 | 270 | 275 | 270.0317
| Std Dev. |  | 0.8165 | 2.0041 | 1.5892 | 3.2042 | 2.6394 | 2.0656 | 1.3784 | 3.4881 | 5.6095 | 2.6583 | 3.2159 | 7.7395 | 21.1048 | 3.3912 | 5.4406 | 5.3072 | 4.6043 | 2.7568 | 4.9187 | 2.2583 | 5.2313 | 1.2247 | 16.76
| Rel.Std.Dev. |  | 0.2998% | 0.7168% | 0.5765% | 1.164% | 0.9879% | 0.7529% | 0.5077% | 1.369% | 2.122% | 0.8901% | 1.304% | 2.789% | 6.272% | 1.226% | 1.957% | 1.979% | 1.737% | 0.9811% | 1.502% | 0.8442% | 1.944% | 0.4446% | 6.31%
| PDM3 |  | 0.1446% | 2.816% | 1.377% | 1.248% | -1.755% | 0.8801% | -0.1618% | -6.291% | -2.797% | 9.828% | -9.328% | 2.045% | 23.73% | 1.677% | 2.228% | -1.388% | -2.552% | 3.332% | 20.44% | -1.633% | -1.02% | 1.309% | -2.325%
| Z-Score (Absolute) |  | 0.22 | 0.58 | 0.15 | 0.11 | 0.78 | 0 | 0.31 | 2.13 | 1.09 | 2.66 | 3.04 | 0.35 | 6.8 | 0.24 | 0.4 | 0.67 | 1.02 | 0.73 | 5.82 | 0.75 | 0.56 | 0.13 | 0.95
| Analytical results for Zr in OREAS 931b (Certified Value 10.8 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 23 |  |  |  | ppm
|  |  | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 0.5g | 01g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.5g | 01g | 0.5g | 0.5g |  | 0.25g | 01g | 0.25g |  |  |  | 1
| 1 | 1 | 13 | 9.8022 | 11.6 | 7.2 | 11.9 | 10 | 11.5 | 10.6 | 10.4 | 11.8 | 10.67 | 8.3 | 12.2 | 12.3 | 16.5 | 12.6 | 8.9199 | 9.8 | 10 | 10.6 |  |  |  | 1
| 1 | 2 | 13 | 9.8398 | 11.42 | 6.6 | 11.5 | 10 | 11.5 | 10.5 | 11.3 | 11.3 | 10.33 | 8.6 | 12.1 | 12.1 | 16.2 | 13.1 | 3.8672 | 9.6 | 11 | 10.2 |  |  |  | 19
| 1 | 3 | 13 | 9.6771 | 11.37 | 6.7 | 11.6 | 10.4 | 11.7 | 10.2 | 10.7 | 11.5 | 10.57 | 8.4 | 12.2 | 12.3 | 15.8 | 13.1 | 7.7319 | 9.8 | 10 | 10.6 |  |  |  | 16
| 1 | 4 | 13 | 9.7853 | 11.32 | 6.8 | 11.9 | 10.1 | 11.5 | 10.6 | 10.7 | 11.2 | 10.67 | 8.6 | 12.5 | 12.2 | 16 | 12.6 | 7.7835 | 10.1 | 10 | 10.4 |  |  |  | 10.7642
| 1 | 5 | 13 | 9.5444 | 11.41 | 6.9 | 11.7 | 10.2 | 11.8 | 10.4 | 10.9 | 11.5 | 10.57 | 7.9 | 12.5 | 12.3 | 16.1 | 12.2 | 8.7852 | 10 | 10 | 11.4 |  |  |  | 119
| 1 | 6 | 13 | 9.4539 | 11.3 | 6.8 | 11.4 | 10.4 | 10.8 | 10.2 | 10.4 | 11.7 | 10.65 | 8.7 | 12.1 | 12.4 | 16.7 | 12.6 | 8.7646 | 9.8 | 10 | 10.2
| Mean |  | 13 | 9.6838 | 11.4033 | 6.8333 | 11.6667 | 10.1833 | 11.4667 | 10.4167 | 10.7333 | 11.5 | 10.5767 | 8.4167 | 12.2667 | 12.2667 | 16.2167 | 12.7 | 7.642 | 9.85 | 10.1667 | 10.5667
| Median |  | 13 | 9.7312 | 11.39 | 6.8 | 11.65 | 10.15 | 11.5 | 10.45 | 10.7 | 11.5 | 10.61 | 8.5 | 12.2 | 12.3 | 16.15 | 12.6 | 8.274 | 9.8 | 10 | 10.5
| Std Dev. |  | 0 | 0.15557 | 0.10746 | 0.20656 | 0.20656 | 0.18348 | 0.35024 | 0.18348 | 0.33862 | 0.22804 | 0.12941 | 0.29269 | 0.18619 | 0.10328 | 0.33116 | 0.34641 | 1.9224 | 0.17607 | 0.40825 | 0.44572
| Rel.Std.Dev. |  | 0% | 1.607% | 0.9423% | 3.023% | 1.771% | 1.802% | 3.054% | 1.761% | 3.155% | 1.983% | 1.224% | 3.477% | 1.518% | 0.842% | 2.042% | 2.728% | 25.16% | 1.787% | 4.016% | 4.218%
| PDM3 |  | 20.77% | -10.04% | 5.937% | -36.52% | 8.384% | -5.396% | 6.526% | -3.229% | -0.2869% | 6.835% | -1.742% | -21.81% | 13.96% | 13.96% | 50.65% | 17.98% | -29.01% | -8.493% | -5.551% | -1.835%
| Z-Score (Absolute) |  | NA | 0.72 | 0.56 | 2.84 | 0.75 | 0.35 | 0.6 | 0.18 | 0.06 | 0.63 | 0.06 | 1.66 | 1.2 | 1.2 | 4.13 | 1.52 | 2.24 | 0.6 | NA | 0.07
Results from laboratories 1 and 22 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C360: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C397: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C525: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C543: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C561: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C579: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C597: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C616: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C634: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C652: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C670: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C707: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C725: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C743: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C761: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C779: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C797: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C815: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C833: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C926: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C945: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C963: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C982: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1001: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1019: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1037: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1074: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1093: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1149: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1185: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 931b (Certified Value 10.81 wt.%) | [15 empty] | Certified
Al2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 10.695 | 10.83 | 11.07 | 10.75 | 10.95 | 10.75 | 10.8 | 11.3 | 9.19 | 11 | 10.81 | 10.71 | 10.6307 | 9.5824 | 1
| 1 | 2 | 10.676 | 10.87 | 11.01 | 10.8 | 10.8 | 11 | 10.8 | 10.8 | 9.203 | 10.6 | 10.79 | 10.78 | 10.5342 | 9.5874 | 0
| 1 | 3 | 10.638 | 10.87 | 10.99 | 10.8 | 10.9 | 11 | 10.75 | 10.6 | 9.203 | 12.4 | 10.74 | 10.72 | 10.6429 | 9.6857 | 16
| 1 | 4 | 10.581 | 10.78 | 11.11 | 10.8 | 10.65 | 10.85 | 10.8 | 10.9 | 9.107 | 11.3 | 10.77 | 10.81 | 10.5255 | 9.6807 | 10.8099
| 1 | 5 | 10.6 | 10.79 | 11.07 | 10.8 | 10.8 | 10.95 | 10.85 | 11.1 | 9.199 | 11.2 | 10.74 | 10.8 | 10.8161 | 9.6639 | 121
| 1 | 6 | 10.713 | 10.82 | 11.16 | 10.75 | 10.85 | 11 | 10.85 | 10.8 | 9.283 | 11.5 | 10.71 | 10.7 | 10.8586 | 9.6662
| Mean |  | 10.6505 | 10.8267 | 11.0683 | 10.7833 | 10.825 | 10.925 | 10.8083 | 10.9167 | 9.1975 | 11.3333 | 10.76 | 10.7533 | 10.668 | 9.6444
| Median |  | 10.657 | 10.825 | 11.07 | 10.8 | 10.825 | 10.975 | 10.8 | 10.85 | 9.201 | 11.25 | 10.755 | 10.75 | 10.6368 | 9.665
| Std Dev. |  | 0.053054 | 0.038297 | 0.062743 | 0.02582 | 0.10368 | 0.10368 | 0.037639 | 0.24833 | 0.055892 | 0.60553 | 0.036878 | 0.048854 | 0.14034 | 0.046822
| Rel.Std.Dev. |  | 0.4981% | 0.3537% | 0.5669% | 0.2394% | 0.9578% | 0.949% | 0.3482% | 2.275% | 0.6077% | 5.343% | 0.3427% | 0.4543% | 1.316% | 0.4855%
| PDM3 |  | -1.474% | 0.1555% | 2.391% | -0.2454% | 0.1401% | 1.065% | -0.01413% | 0.988% | -14.92% | 4.843% | -0.4612% | -0.5229% | -1.312% | -10.78%
| Z-Score (Absolute) |  | 0.79 | 0.17 | 1.48 | 0.07 | 0.16 | 0.7 | 0.07 | 0.66 | 8.67 | 2.91 | 0.19 | 0.23 | 0.69 | 6.24
| Analytical results for BaO in OREAS 931b (Certified Value 2009 ppm) | [15 empty] | Certified
BaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 11 | 12 | 14 | 16 | 19 | 24 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF |  | 0
0
| 1 | 1 | 2289 | 2000 | 2620 | 1900 | 2000 | 1900 | 1900 | 2000 | 2300 | 2200 | 1900 | 1943.3 | 1890.34 |  | 1
| 1 | 2 | 2222 | 2200 | 2690 | 1900 | 1900 | 2100 | 1800 | 2010 | 1500 | 2100 | 1900 | 1953.5 | 1950.26 |  | 0
| 1 | 3 | 2322 | 2200 | 2610 | 1900 | 1900 | 2000 | 1800 | 2020 | 1800 | 2100 | 1900 | 1903.25 | 1905.343 |  | 16
| 1 | 4 | 2367 | 2100 | 2660 | 2000 | 1800 | 2000 | 1900 | 1970 | 2100 | 2100 | 1900 | 1969.25 | 1959.363 |  | 2008.6475
| 1 | 5 | 2255 | 2100 | 2790 | 1900 | 1900 | 2000 | 1800 | 2010 | 2300 | 2000 | 1790 | 2027.1 | 1954.512 |  | 122
| 1 | 6 | 2311 | 2100 | 2870 | 1900 | 1900 | 2000 | 1900 | 2050 | 2200 | 2100 | 1900 | 2051 | 1872.306
| Mean |  | 2294.3333 | 2116.6667 | 2706.6667 | 1916.6667 | 1900 | 2000 | 1850 | 2010 | 2033.3333 | 2100 | 1881.6667 | 1974.5667 | 1922.0207
| Median |  | 2300 | 2100 | 2675 | 1900 | 1900 | 2000 | 1850 | 2010 | 2150 | 2100 | 1900 | 1961.375 | 1927.8015
| Std Dev. |  | 51.2471 | 75.2773 | 102.8915 | 40.8248 | 63.2456 | 63.2456 | 54.7723 | 26.0768 | 320.4164 | 63.2456 | 44.9073 | 55.0228 | 37.4191
| Rel.Std.Dev. |  | 2.234% | 3.556% | 3.801% | 2.13% | 3.329% | 3.162% | 2.961% | 1.297% | 15.76% | 3.012% | 2.387% | 2.787% | 1.947%
| PDM3 |  | 14.22% | 5.378% | 34.75% | -4.579% | -5.409% | -0.4305% | -7.898% | 0.06733% | 1.229% | 4.548% | -6.322% | -1.697% | -4.313%
| Z-Score (Absolute) |  | 1.99 | 0.79 | 4.77 | 0.56 | 0.67 | 0 | 1.01 | 0.07 | 0.22 | 0.67 | 0.81 | 0.17 | 0.53
| Analytical results for Bi in OREAS 931b (Indicative Value 212 ppm) | [15 empty] | Indicative
Bi | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 207.4447 | [13 empty] | 1
| 1 | 2 | 217.379 | [13 empty] | 8
| 1 | 3 | 219.3488 | [13 empty] | 16
| 1 | 4 | 195.2832 | [13 empty] | 212.0613
| 1 | 5 | 214.6772 | [13 empty] | 14
| 1 | 6 | 218.2347
| Mean |  | 212.0613
| Median |  | 216.0281
| Std Dev. |  | 9.2645
| Rel.Std.Dev. |  | 4.369%
| PDM3 |  | 7.859e-07%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 931b (Certified Value 0.115 wt.%) | [15 empty] | Certified
CaO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 3
3
| 1 | 1 | 0.1259 | 0.13 | 0.06 | 0.12 | 0.13 | 0.1 | 0.12 | 0.11 | 0.108 | 0.11 | 0.12 | 0.12 | 0.1708 | 0.096528 | 1
| 1 | 2 | 0.1259 | 0.13 | 0.07 | 0.11 | 0.13 | 0.1 | 0.12 | 0.07 | 0.107 | 0.11 | 0.12 | 0.12 | 0.1693 | 0.084242 | 0
| 1 | 3 | 0.1259 | 0.13 | 0.07 | 0.11 | 0.11 | 0.1 | 0.12 | 0.08 | 0.101 | 0.33 | 0.12 | 0.12 | 0.17245 | 0.10698 | 16
| 1 | 4 | 0.1119 | 0.13 | 0.07 | 0.11 | 0.12 | 0.1 | 0.12 | 0.07 | 0.117 | 0.12 | 0.12 | 0.12 | 0.16115 | 0.10056 | 0.11503
| 1 | 5 | 0.1119 | 0.13 | 0.07 | 0.11 | 0.12 | 0.1 | 0.11 | 0.09 | 0.113 | 0.12 | 0.13 | 0.12 | 0.16898 | 0.1057 | 123
| 1 | 6 | 0.1259 | 0.13 | 0.07 | 0.11 | 0.12 | 0.1 | 0.11 | 0.07 | 0.104 | 0.12 | 0.12 | 0.12 | 0.16889 | 0.094144
| Mean |  | 0.12123 | 0.13 | 0.068333 | 0.11167 | 0.12167 | 0.1 | 0.11667 | 0.081667 | 0.10833 | 0.15167 | 0.12167 | 0.12 | 0.16859 | 0.098027
| Median |  | 0.1259 | 0.13 | 0.07 | 0.11 | 0.12 | 0.1 | 0.12 | 0.075 | 0.1075 | 0.12 | 0.12 | 0.12 | 0.16914 | 0.098544
| Std Dev. |  | 0.0072296 | 0 | 0.0040825 | 0.0040825 | 0.0075277 | 1.5202e-17 | 0.005164 | 0.016021 | 0.0058538 | 0.087502 | 0.0040825 | 0 | 0.0038945 | 0.0084015
| Rel.Std.Dev. |  | 5.963% | 0% | 5.974% | 3.656% | 6.187% | 1.52e-14% | 4.426% | 19.62% | 5.403% | 57.69% | 3.355% | 0% | 2.31% | 8.571%
| PDM3 |  | 5.396% | 13.02% | -40.59% | -2.921% | 5.773% | -13.06% | 1.426% | -29% | -5.819% | 31.85% | 5.773% | 4.324% | 46.57% | -14.78%
| Z-Score (Absolute) |  | 0.18 | 0.73 | 3.11 | 0.41 | 0.21 | 1.14 | 0.1 | 2.28 | 0.62 | 2.07 | 0.21 | 0.1 | 3.13 | 1.26
| Analytical results for Ce in OREAS 931b (Indicative Value < 80 ppm) | [15 empty] | Indicative
Ce | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 81.85 | [13 empty] | 1
| 1 | 2 | < 150 | [13 empty] | 9
| 1 | 3 | < 80 | [13 empty] | 16
| 1 | 4 | < 80 | [13 empty] | < 80
| 1 | 5 | < 80 | [13 empty] | 15
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 931b (Indicative Value < 331.815 ppm) | [15 empty] | Indicative
Cl | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 332 | [13 empty] | 1
| 1 | 2 | < 439 | [13 empty] | 10
| 1 | 3 | < 326 | [13 empty] | 16
| 1 | 4 | < 384 | [13 empty] | < 331.815
| 1 | 5 | < 289 | [13 empty] | 16
| 1 | 6 | < 349
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 931b (Indicative Value < 60 ppm) | [15 empty] | Indicative
Co | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 19 | 24 | [12 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | < 60 | < 39 | [12 empty] | 1
| 1 | 2 | < 60 | < 39 | [12 empty] | 11
| 1 | 3 | < 60 | 40.0674 | [12 empty] | 16
| 1 | 4 | < 60 | 49.7362 | [12 empty] | < 60
| 1 | 5 | < 60 | 46.2726 | [12 empty] | 17
| 1 | 6 | < 60 | < 39
| Mean |  |  | 45.3587
| Median |  |  | 46.2726
| Std Dev. |  |  | 4.8987
| Rel.Std.Dev. |  |  | 10.8%
PDM3
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cr2O3 in OREAS 931b (Indicative Value 129 ppm) | [15 empty] | Indicative
Cr2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 16 | 19 | 24 | [9 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 0
0
| 1 | 1 | 100 | 200 | < 100 | < 160 | < 50 | [9 empty] | 1
| 1 | 2 | 200 | 200 | < 100 | < 160 | 69.24 | [9 empty] | 12
| 1 | 3 | 100 | 200 | < 100 | < 160 | 91.701 | [9 empty] | 16
| 1 | 4 | 100 | 100 | < 100 | < 160 | < 50 | [9 empty] | 128.9203
| 1 | 5 | 300 | 200 | < 100 | < 160 | < 50 | [9 empty] | 18
| 1 | 6 | 200 | 100 | < 100 | < 160 | 84.624
| Mean |  | 166.6667 | 166.6667 |  |  | 81.855
| Median |  | 150 | 200 |  |  | 84.624
| Std Dev. |  | 81.6497 | 51.6398 |  |  | 11.4837
| Rel.Std.Dev. |  | 48.99% | 30.98% |  |  | 14.03%
| PDM3 |  | 29.28% | 29.28% |  |  | -36.51%
| Z-Score (Absolute) |  | 1.95 | 1.95 | 0.67 | 0 | 0.6
| Analytical results for Cu in OREAS 931b (Certified Value 3.95 wt.%) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 4.087 | 3.929 | > 3 | 3.84 | 4.02 | 3.91 | 3.83 | 4.01 | 4.3637 | 4.12 | 4.01 | 3.9 | 3.776 | 3.6129 | 1
| 1 | 2 | 4.101 | 3.947 | > 3 | 3.86 | 3.96 | 4.01 | 3.83 | 3.91 | 4.3469 | 3.84 | 3.986 | 3.94 | 3.7631 | 3.5666 | 0
| 1 | 3 | 4.106 | 3.931 | > 3 | 3.83 | 4 | 3.94 | 3.82 | 3.9 | 4.2806 | 4.16 | 3.994 | 3.92 | 3.8035 | 3.6246 | 16
| 1 | 4 | 4.106 | 3.958 | > 3 | 3.86 | 3.98 | 3.91 | 3.81 | 4.04 | 4.299 | 4.26 | 4.018 | 3.96 | 3.7361 | 3.6295 | 3.9506
| 1 | 5 | 4.065 | 3.96 | > 3 | 3.89 | 4.03 | 3.94 | 3.81 | 3.99 | 4.3006 | 4.23 | 3.994 | 3.98 | 3.8314 | 3.6135 | 124
| 1 | 6 | 4.088 | 3.954 | > 3 | 3.89 | 3.99 | 3.92 | 3.89 | 3.94 | 4.3045 | 4.37 | 3.994 | 3.92 | 3.8436 | 3.6063
| Mean |  | 4.0922 | 3.9465 |  | 3.8617 | 3.9967 | 3.9383 | 3.8317 | 3.965 | 4.3159 | 4.1633 | 3.9993 | 3.9367 | 3.7923 | 3.6089
| Median |  | 4.0945 | 3.9505 |  | 3.86 | 3.995 | 3.93 | 3.825 | 3.965 | 4.3025 | 4.195 | 3.994 | 3.93 | 3.7898 | 3.6132
| Std Dev. |  | 0.015766 | 0.013546 |  | 0.024833 | 0.02582 | 0.037639 | 0.029944 | 0.056833 | 0.032066 | 0.18052 | 0.012044 | 0.029439 | 0.041389 | 0.022371
| Rel.Std.Dev. |  | 0.3853% | 0.3432% |  | 0.6431% | 0.646% | 0.9557% | 0.7815% | 1.433% | 0.743% | 4.336% | 0.3012% | 0.7478% | 1.091% | 0.6199%
| PDM3 |  | 3.584% | -0.1036% |  | -2.251% | 1.166% | -0.3103% | -3.01% | 0.3647% | 9.246% | 5.385% | 1.234% | -0.3525% | -4.007% | -8.649%
| Z-Score (Absolute) |  | 1.16 | 0 | NA | 0.67 | 0.4 | 0.06 | 0.91 | 0.15 | 2.94 | 1.72 | 0.42 | 0.08 | 1.23 | 2.68
| Analytical results for F in OREAS 931b (Indicative Value < 3000 ppm) | [15 empty] | Indicative
F | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 3000 | [13 empty] | 1
| 1 | 2 | < 3000 | [13 empty] | 13
| 1 | 3 | < 3000 | [13 empty] | 16
| 1 | 4 | < 3000 | [13 empty] | < 3000
| 1 | 5 | < 3000 | [13 empty] | 19
| 1 | 6 | < 3000
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 931b (Certified Value 12.42 wt.%) | [15 empty] | Certified
Fe2O3 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 12.51 | 12.43 | 12.538 | 12.25 | 12.65 | 12.3 | 12.35 | 12.3 | 13.13 | 12.481 | 12.49 | 12.5 | 12.2368 | 11.7473 | 1
| 1 | 2 | 12.39 | 12.45 | 12.596 | 12.25 | 12.55 | 12.55 | 12.3 | 12 | 13.119 | 11.924 | 12.43 | 12.46 | 12.179 | 11.7177 | 0
| 1 | 3 | 12.34 | 12.41 | 12.553 | 12.2 | 12.65 | 12.5 | 12.3 | 11.9 | 12.993 | 12.967 | 12.48 | 12.44 | 12.3305 | 11.7683 | 16
| 1 | 4 | 12.49 | 12.48 | 12.41 | 12.35 | 12.55 | 12.4 | 12.35 | 12.2 | 12.883 | 13.153 | 12.5 | 12.46 | 12.1614 | 11.7894 | 12.4208
| 1 | 5 | 12.52 | 12.46 | 12.438 | 12.4 | 12.7 | 12.45 | 12.3 | 12.2 | 13.044 | 13.082 | 12.49 | 12.52 | 12.4578 | 11.6832 | 125
| 1 | 6 | 12.5 | 12.49 | 12.438 | 12.3 | 12.6 | 12.4 | 12.4 | 12 | 13.038 | 13.411 | 12.45 | 12.49 | 12.4444 | 11.7162
| Mean |  | 12.4583 | 12.4533 | 12.4955 | 12.2917 | 12.6167 | 12.4333 | 12.3333 | 12.1 | 13.0345 | 12.8363 | 12.4733 | 12.4783 | 12.3017 | 11.737
| Median |  | 12.495 | 12.455 | 12.488 | 12.275 | 12.625 | 12.425 | 12.325 | 12.1 | 13.041 | 13.0245 | 12.485 | 12.475 | 12.2837 | 11.7325
| Std Dev. |  | 0.074677 | 0.030111 | 0.076335 | 0.073598 | 0.060553 | 0.08756 | 0.040825 | 0.15492 | 0.090595 | 0.54158 | 0.027325 | 0.029944 | 0.13 | 0.038795
| Rel.Std.Dev. |  | 0.5994% | 0.2418% | 0.6109% | 0.5988% | 0.4799% | 0.7042% | 0.331% | 1.28% | 0.695% | 4.219% | 0.2191% | 0.24% | 1.057% | 0.3305%
| PDM3 |  | 0.3025% | 0.2623% | 0.6018% | -1.039% | 1.577% | 0.1012% | -0.7039% | -2.582% | 4.941% | 3.346% | 0.4233% | 0.4635% | -0.9589% | -5.505%
| Z-Score (Absolute) |  | 0.01 | 0.01 | 0.19 | 0.8 | 0.78 | 0.11 | 0.6 | 1.73 | 2.82 | 1.85 | 0.09 | 0.11 | 0.75 | 3.5
| Analytical results for Ga in OREAS 931b (Indicative Value < 70 ppm) | [15 empty] | Indicative
Ga | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 70 | [13 empty] | 1
| 1 | 2 | < 70 | [13 empty] | 14
| 1 | 3 | < 70 | [13 empty] | 16
| 1 | 4 | < 70 | [13 empty] | < 70
| 1 | 5 | < 70 | [13 empty] | 20
| 1 | 6 | < 70
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 931b (Indicative Value < 90 ppm) | [15 empty] | Indicative
Gd | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 90 | [13 empty] | 1
| 1 | 2 | < 90 | [13 empty] | 15
| 1 | 3 | < 90 | [13 empty] | 16
| 1 | 4 | < 90 | [13 empty] | < 90
| 1 | 5 | < 90 | [13 empty] | 21
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 931b (Indicative Value < 80 ppm) | [15 empty] | Indicative
Hf | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 80 | [13 empty] | 1
| 1 | 2 | < 80 | [13 empty] | 16
| 1 | 3 | < 80 | [13 empty] | 16
| 1 | 4 | < 80 | [13 empty] | < 80
| 1 | 5 | < 80 | [13 empty] | 22
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 931b (Certified Value 2.55 wt.%) | [15 empty] | Certified
K2O | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 16 | 19 | 24 | [9 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 2
3
| 1 | 1 | 2.59 | 2.69 | 2.59 | 2.4257 | 2.4514 | [9 empty] | 1
| 1 | 2 | 2.58 | 2.69 | 2.63 | 2.4233 | 2.4315 | [9 empty] | 0
| 1 | 3 | 2.55 | 2.67 | 2.62 | 2.4174 | 2.4678 | [9 empty] | 16
| 1 | 4 | 2.58 | 2.7 | 2.62 | 2.4343 | 2.458 | [9 empty] | 2.5527
| 1 | 5 | 2.6 | 2.69 | 2.64 | 2.4145 | 2.4581 | [9 empty] | 126
| 1 | 6 | 2.58 | 2.69 | 2.61 | 2.4338 | 2.4467
| Mean |  | 2.58 | 2.6883 | 2.6183 | 2.4248 | 2.4522
| Median |  | 2.58 | 2.69 | 2.62 | 2.4245 | 2.4547
| Std Dev. |  | 0.016733 | 0.0098319 | 0.017224 | 0.0081762 | 0.012422
| Rel.Std.Dev. |  | 0.6486% | 0.3657% | 0.6578% | 0.3372% | 0.5066%
| PDM3 |  | 1.068% | 5.311% | 2.569% | -5.011% | -3.937%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.24 | 0.97 | 0.8
| Analytical results for La in OREAS 931b (Indicative Value < 170 ppm) | [15 empty] | Indicative
La | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 170 | [13 empty] | 1
| 1 | 2 | < 170 | [13 empty] | 17
| 1 | 3 | < 170 | [13 empty] | 16
| 1 | 4 | < 170 | [13 empty] | < 170
| 1 | 5 | < 170 | [13 empty] | 23
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 931b (Certified Value 0.632 wt.%) | [15 empty] | Certified
MgO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 16 | 19 | 24 | [9 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 3
3
| 1 | 1 | 0.68 | 0.57 | 0.64 | 0.70844 | 0.55705 | [9 empty] | 1
| 1 | 2 | 0.72 | 0.58 | 0.63 | 0.75579 | 0.5493 | [9 empty] | 0
| 1 | 3 | 0.6 | 0.58 | 0.62 | 0.7416 | 0.57466 | [9 empty] | 16
| 1 | 4 | 0.57 | 0.57 | 0.62 | 0.73242 | 0.5733 | [9 empty] | 0.63185
| 1 | 5 | 0.59 | 0.61 | 0.63 | 0.78443 | 0.58935 | [9 empty] | 127
| 1 | 6 | 0.62 | 0.61 | 0.63 | 0.74176 | 0.57756
| Mean |  | 0.63 | 0.58667 | 0.62833 | 0.74407 | 0.5702
| Median |  | 0.61 | 0.58 | 0.63 | 0.74168 | 0.57398
| Std Dev. |  | 0.057966 | 0.018619 | 0.0075277 | 0.025232 | 0.014559
| Rel.Std.Dev. |  | 9.201% | 3.174% | 1.198% | 3.391% | 2.553%
| PDM3 |  | -0.2936% | -7.152% | -0.5573% | 17.76% | -9.757%
| Z-Score (Absolute) |  | 0.03 | 0.67 | 0 | 1.87 | 0.94
| Analytical results for MnO in OREAS 931b (Certified Value 0.022 wt.%) | [15 empty] | Certified
MnO | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 16 | 19 | 24 | [9 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 3
3
| 1 | 1 | 0.03 | 0.02 | 0.03 | 0.019949 | 0.017094 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.02 | 0.02 | 0.022216 | 0.01731 | [9 empty] | 0
| 1 | 3 | 0.03 | 0.02 | 0.02 | 0.015979 | 0.019359 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.02 | 0.02 | 0.019807 | 0.015551 | [9 empty] | 0.02226
| 1 | 5 | 0.03 | 0.02 | 0.03 | 0.02013 | 0.014958 | [9 empty] | 128
| 1 | 6 | 0.03 | 0.02 | 0.03 | 0.021699 | 0.013751
| Mean |  | 0.03 | 0.02 | 0.025 | 0.019963 | 0.016337
| Median |  | 0.03 | 0.02 | 0.025 | 0.02004 | 0.016322
| Std Dev. |  | 0 | 0 | 0.0054772 | 0.0021915 | 0.0019929
| Rel.Std.Dev. |  | 0% | 0% | 21.91% | 10.98% | 12.2%
| PDM3 |  | 34.77% | -10.15% | 12.31% | -10.32% | -26.61%
| Z-Score (Absolute) |  | 1.84 | 0 | 0.92 | 0.01 | 0.67
| Analytical results for Mo in OREAS 931b (Indicative Value < 33 ppm) | [15 empty] | Indicative
Mo | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 1
1
| 1 | 1 | < 33 | [13 empty] | 1
| 1 | 2 | < 33 | [13 empty] | 8
| 1 | 3 | < 33 | [13 empty] | 16
| 1 | 4 | < 33 | [13 empty] | < 33
| 1 | 5 | < 33 | [13 empty] | 14
| 1 | 6 | < 33
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 931b (Indicative Value 0.154 wt.%) | [15 empty] | Indicative
Na2O | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 01 | 24 | [12 empty] | wt.%
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 3
3
| 1 | 1 | 0.37 | 0.12267 | [12 empty] | 1
| 1 | 2 | 0.61 | 0.043852 | [12 empty] | 9
| 1 | 3 | 0.22 | 0.077436 | [12 empty] | 16
| 1 | 4 | 0.2 | 0.068422 | [12 empty] | 0.15426
| 1 | 5 | 0.18 | 0.080773 | [12 empty] | 15
| 1 | 6 | 0.2 | 0.053948
| Mean |  | 0.29667 | 0.074517
| Median |  | 0.21 | 0.072929
| Std Dev. |  | 0.16836 | 0.027448
| Rel.Std.Dev. |  | 56.75% | 36.83%
| PDM3 |  | 92.32% | -51.69%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Nb in OREAS 931b (Indicative Value < 59.95051 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 55.53 | [13 empty] | 1
| 1 | 2 | < 35 | [13 empty] | 10
| 1 | 3 | < 64.1 | [13 empty] | 16
| 1 | 4 | < 62.33 | [13 empty] | < 59.95051
| 1 | 5 | < 59.95 | [13 empty] | 16
| 1 | 6 | < 99.09
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 931b (Indicative Value < 90 ppm) | [15 empty] | Indicative
Nd | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 90 | [13 empty] | 1
| 1 | 2 | < 90 | [13 empty] | 5
| 1 | 3 | < 90 | [13 empty] | 16
| 1 | 4 | < 90 | [13 empty] | < 90
| 1 | 5 | < 90 | [13 empty] | 17
| 1 | 6 | < 90
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 931b (Indicative Value 201 ppm) | [15 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 19 | 24 | [12 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 176.35 | < 80 | [12 empty] | 1
| 1 | 2 | 192.9 | < 80 | [12 empty] | 12
| 1 | 3 | 198.3 | < 80 | [12 empty] | 16
| 1 | 4 | 209.55 | < 80 | [12 empty] | 201.0917
| 1 | 5 | 219.65 | < 80 | [12 empty] | 18
| 1 | 6 | 209.8 | < 80
| Mean |  | 201.0917
| Median |  | 203.925
| Std Dev. |  | 15.3547
| Rel.Std.Dev. |  | 7.636%
| PDM3 |  | -1.665e-13%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for P2O5 in OREAS 931b (Certified Value 0.07 wt.%) | [15 empty] | Certified
P2O5 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 16 | 19 | 24 | [9 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 3
3
| 1 | 1 | 0.08 | 0.073 | 0.07 | 0.089693 | 0.045248 | [9 empty] | 1
| 1 | 2 | 0.08 | 0.067 | 0.07 | 0.103 | 0.039236 | [9 empty] | 0
| 1 | 3 | 0.08 | 0.067 | 0.07 | 0.081512 | 0.047888 | [9 empty] | 16
| 1 | 4 | 0.07 | 0.07 | 0.07 | 0.091526 | 0.045615 | [9 empty] | 0.069862
| 1 | 5 | 0.07 | 0.072 | 0.07 | 0.087321 | 0.045871 | [9 empty] | 129
| 1 | 6 | 0.08 | 0.064 | 0.07 | 0.088891 | 0.045485
| Mean |  | 0.076667 | 0.068833 | 0.07 | 0.090323 | 0.044891
| Median |  | 0.08 | 0.0685 | 0.07 | 0.089292 | 0.04555
| Std Dev. |  | 0.005164 | 0.0034303 | 0 | 0.0070872 | 0.0029301
| Rel.Std.Dev. |  | 6.736% | 4.983% | 0% | 7.846% | 6.527%
| PDM3 |  | 9.74% | -1.473% | 0.1974% | 29.29% | -35.74%
| Z-Score (Absolute) |  | 0.67 | 0.12 | 0 | 2.06 | 2.54
| Analytical results for Pb in OREAS 931b (Indicative Value 136 ppm) | [15 empty] | Indicative
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 0
0
| 1 | 1 | 60 | 120 | 310 | 110 | 120 | 50 | 150 | 100 | < 200 | 1500 | 190 | 200 | < 100 | 186.6877 | 1
| 1 | 2 | 40 | 170 | 310 | 100 | 120 | 90 | 120 | < 100 | < 200 | 1500 | 190 | 100 | < 100 | 203.5462 | 13
| 1 | 3 | 30 | 150 | 270 | 110 | 90 | 70 | 130 | < 100 | < 200 | < 100 | 190 | 100 | < 100 | 189.1756 | 16
| 1 | 4 | 80 | 130 | 280 | 120 | < 50 | 70 | 140 | 200 | < 200 | < 100 | 190 | 200 | < 100 | 182.8565 | 136.2155
| 1 | 5 | 50 | 160 | 250 | 110 | 90 | 80 | 130 | 100 | < 200 | < 100 | 190 | 200 | < 100 | 175.8893 | 19
| 1 | 6 | 50 | 140 | 200 | 100 | 70 | 50 | 160 | < 100 | < 200 | < 100 | 190 | 100 | 100.1 | 176.7583
| Mean |  | 51.6667 | 145 | 270 | 108.3333 | 98 | 68.3333 | 138.3333 | 133.3333 |  | 1500 | 190 | 150 | 100.1 | 185.8189
| Median |  | 50 | 145 | 275 | 110 | 90 | 70 | 135 | 100 |  | 1500 | 190 | 150 | 100.1 | 184.7721
| Std Dev. |  | 17.224 | 18.7083 | 41.4729 | 7.5277 | 21.6795 | 16.0208 | 14.7196 | 57.735 |  | 0 | 0 | 54.7723 |  | 10.1562
| Rel.Std.Dev. |  | 33.34% | 12.9% | 15.36% | 6.949% | 22.12% | 23.45% | 10.64% | 43.3% |  | 0% | 0% | 36.51% |  | 5.466%
| PDM3 |  | -62.07% | 6.449% | 98.22% | -20.47% | -28.06% | -49.83% | 1.555% | -2.116% |  | 1001% | 39.48% | 10.12% | -26.51% | 36.42%
| Z-Score (Absolute) |  | 1.04 | 0.32 | 2.14 | 0.22 | 0.55 | 0.8 | 0.22 | 0.46 | 0.34 | 5.98 | 0.91 | 0.39 | 0.95 | 0.91
| Analytical results for Pr in OREAS 931b (Indicative Value < 80 ppm) | [15 empty] | Indicative
Pr | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 80 | [13 empty] | 1
| 1 | 2 | < 80 | [13 empty] | 14
| 1 | 3 | < 80 | [13 empty] | 16
| 1 | 4 | < 80 | [13 empty] | < 80
| 1 | 5 | < 80 | [13 empty] | 20
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 931b (Indicative Value 140 ppm) | [15 empty] | Indicative
Rb | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | 143.4327 | [13 empty] | 1
| 1 | 2 | 138.2352 | [13 empty] | 15
| 1 | 3 | 141.6174 | [13 empty] | 16
| 1 | 4 | 145.0394 | [13 empty] | 139.8011
| 1 | 5 | 133.1302 | [13 empty] | 21
| 1 | 6 | 137.3515
| Mean |  | 139.8011
| Median |  | 139.9263
| Std Dev. |  | 4.4018
| Rel.Std.Dev. |  | 3.149%
| PDM3 |  | 9.942e-08%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 931b (Certified Value 4.35 wt.%) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 19 | 24 |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF |  | 2
3
| 1 | 1 | 4.58 | 4.537 | > 2 | 4.55 | 4.58 | 4.33 | 4.47 | 4.27 | 3.2272 | 4.31 | 4.445 | 3.9621 | 4.0543 |  | 1
| 1 | 2 | 4.56 | 4.517 | > 2 | 4.54 | 4.49 | 4.52 | 4.43 | 4.1 | 3.1736 | 3.98 | 4.445 | 4.0022 | 4.0519 |  | 0
| 1 | 3 | 4.64 | 4.553 | > 2 | 4.52 | 4.56 | 4.45 | 4.44 | 4.04 | 3.256 | 3.83 | 4.445 | 4.1264 | 4.0243 |  | 16
| 1 | 4 | 4.64 | 4.557 | > 2 | 4.54 | 4.52 | 4.42 | 4.43 | 4.21 | 3.3088 | 3.97 | 4.485 | 4.1139 | 4.0307 |  | 4.3509
| 1 | 5 | 4.64 | 4.569 | > 2 | 4.49 | 4.6 | 4.44 | 4.43 | 4.24 | 3.2148 | 3.68 | 4.485 | 4.1163 | 4.022 |  | 130
| 1 | 6 | 4.58 | 4.541 | > 2 | 4.37 | 4.55 | 4.46 | 4.47 | 4.18 | 3.2196 | 4.13 | 4.445 | 4.1279 | 4.0309
| Mean |  | 4.6067 | 4.5457 |  | 4.5017 | 4.55 | 4.4367 | 4.445 | 4.1733 | 3.2333 | 3.9833 | 4.4583 | 4.0748 | 4.0357
| Median |  | 4.61 | 4.547 |  | 4.53 | 4.555 | 4.445 | 4.435 | 4.195 | 3.2234 | 3.975 | 4.445 | 4.1151 | 4.0308
| Std Dev. |  | 0.037238 | 0.01814 |  | 0.067946 | 0.04 | 0.062183 | 0.019748 | 0.08756 | 0.04549 | 0.22088 | 0.020656 | 0.073079 | 0.013946
| Rel.Std.Dev. |  | 0.8083% | 0.3991% |  | 1.509% | 0.8791% | 1.402% | 0.4443% | 2.098% | 1.407% | 5.545% | 0.4633% | 1.793% | 0.3456%
| PDM3 |  | 5.878% | 4.476% |  | 3.465% | 4.576% | 1.971% | 2.162% | -4.082% | -25.69% | -8.449% | 2.469% | -6.346% | -7.245%
| Z-Score (Absolute) |  | 0.81 | 0.51 | NA | 0.3 | 0.54 | 0.02 | 0.02 | 1.31 | 5.92 | 2.24 | 0.09 | 1.8 | 1.99
| Analytical results for Sb in OREAS 931b (Indicative Value < 170 ppm) | [15 empty] | Indicative
Sb | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 170 | [13 empty] | 1
| 1 | 2 | < 170 | [13 empty] | 16
| 1 | 3 | < 170 | [13 empty] | 16
| 1 | 4 | < 170 | [13 empty] | < 170
| 1 | 5 | < 170 | [13 empty] | 22
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 931b (Certified Value 62.36 wt.%) | [15 empty] | Certified
SiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 2
3
| 1 | 1 | 61.99 | 62.74 | 62.63 | 61.5 | 62.4 | 61.5 | 62.1 | 63.8 | 59.711 | 60.6 | 62.49 | 61.79 | 62.3755 | 55.2643 | 1
| 1 | 2 | 61.44 | 62.77 | 62.39 | 62 | 61.8 | 63.3 | 62.1 | 62 | 59.874 | 58.1 | 62.63 | 61.75 | 61.9707 | 54.9131 | 0
| 1 | 3 | 61.37 | 62.54 | 62.45 | 61.9 | 62.3 | 63 | 62 | 61.5 | 59.749 | 64.5 | 62.54 | 61.71 | 62.2545 | 55.4108 | 16
| 1 | 4 | 61.92 | 62.73 | 62.84 | 62.3 | 61.4 | 62.3 | 62.3 | 61.8 | 58.987 | 64.2 | 62.6 | 61.83 | 61.454 | 55.4054 | 62.3625
| 1 | 5 | 61.82 | 62.97 | 62.42 | 62.1 | 62.7 | 62.5 | 61.9 | 63.2 | 59.908 | 63.2 | 62.48 | 61.87 | 62.7506 | 55.4304 | 131
| 1 | 6 | 61.97 | 63.03 | 62.83 | 61.8 | 62.1 | 63 | 62.4 | 62.5 | 59.893 | 64.7 | 62.12 | 61.74 | 62.753 | 55.2722
| Mean |  | 61.7517 | 62.7967 | 62.5933 | 61.9333 | 62.1167 | 62.6 | 62.1333 | 62.4667 | 59.687 | 62.55 | 62.4767 | 61.7817 | 62.2597 | 55.2827
| Median |  | 61.87 | 62.755 | 62.54 | 61.95 | 62.2 | 62.75 | 62.1 | 62.25 | 59.8115 | 63.7 | 62.515 | 61.77 | 62.315 | 55.3388
| Std Dev. |  | 0.27578 | 0.17818 | 0.205 | 0.27325 | 0.46224 | 0.65115 | 0.18619 | 0.88468 | 0.35233 | 2.6524 | 0.18446 | 0.060139 | 0.49607 | 0.19511
| Rel.Std.Dev. |  | 0.4466% | 0.2837% | 0.3275% | 0.4412% | 0.7441% | 1.04% | 0.2997% | 1.416% | 0.5903% | 4.24% | 0.2953% | 0.09734% | 0.7968% | 0.3529%
| PDM3 |  | -0.9795% | 0.6962% | 0.3702% | -0.6881% | -0.3942% | 0.3809% | -0.3674% | 0.1671% | -4.29% | 0.3007% | 0.1831% | -0.9313% | -0.1648% | -11.35%
| Z-Score (Absolute) |  | 0.8 | 1.08 | 0.71 | 0.47 | 0.14 | 0.73 | 0.11 | 0.49 | 4.51 | 0.64 | 0.5 | 0.75 | 0.11 | 12.43
| Analytical results for Sm in OREAS 931b (Indicative Value < 170 ppm) | [15 empty] | Indicative
Sm | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 170 | [13 empty] | 1
| 1 | 2 | < 170 | [13 empty] | 17
| 1 | 3 | < 170 | [13 empty] | 16
| 1 | 4 | < 170 | [13 empty] | < 170
| 1 | 5 | < 170 | [13 empty] | 23
| 1 | 6 | < 170
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 931b (Indicative Value < 160 ppm) | [15 empty] | Indicative
Sn | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 160 | [13 empty] | 1
| 1 | 2 | < 160 | [13 empty] | 8
| 1 | 3 | < 160 | [13 empty] | 16
| 1 | 4 | < 160 | [13 empty] | < 160
| 1 | 5 | < 160 | [13 empty] | 14
| 1 | 6 | < 160
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for SrO in OREAS 931b (Indicative Value 51 ppm) | [15 empty] | Indicative
SrO | Round | Replicate | Lab | Lab | Lab | [11 empty] | 1
| No | No | 01 | 19 | 24 | [11 empty] | ppm
|  |  | BF*XRF | BF*XRF | OxBF*XRF | [11 empty] | 0
0
| 1 | 1 | < 100 | < 30 | 92.506 | [11 empty] | 1
| 1 | 2 | < 100 | 37.7 | 72.702 | [11 empty] | 9
| 1 | 3 | < 100 | < 30 | 85.5876 | [11 empty] | 16
| 1 | 4 | < 100 | < 30 | 66.3488 | [11 empty] | 51.4444
| 1 | 5 | 100 | < 30 | 67.8096 | [11 empty] | 15
| 1 | 6 | < 100 | 37.5 | 74.046
| Mean |  | 100 | 37.6 | 76.5
| Median |  | 100 | 37.6 | 73.374
| Std Dev. |  |  | 0.14142 | 10.371
| Rel.Std.Dev. |  |  | 0.3761% | 13.56%
| PDM3 |  | 94.38% | -26.91% | 48.7%
| Z-Score (Absolute) |  | 0 | 1.33 | 0.67
| Analytical results for Ta in OREAS 931b (Indicative Value < 80 ppm) | [15 empty] | Indicative
Ta | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 80 | [13 empty] | 1
| 1 | 2 | < 80 | [13 empty] | 10
| 1 | 3 | < 80 | [13 empty] | 16
| 1 | 4 | < 80 | [13 empty] | < 80
| 1 | 5 | < 80 | [13 empty] | 16
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 931b (Indicative Value < 44 ppm) | [15 empty] | Indicative
Th | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 1
1
| 1 | 1 | < 44 | [13 empty] | 1
| 1 | 2 | < 44 | [13 empty] | 11
| 1 | 3 | 44.7711 | [13 empty] | 16
| 1 | 4 | < 44 | [13 empty] | < 44
| 1 | 5 | < 44 | [13 empty] | 17
| 1 | 6 | < 44
| Mean |  | 44.7711
| Median |  | 44.7711
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 931b (Certified Value 0.441 wt.%) | [15 empty] | Certified
TiO2 | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 16 | 19 | 24 | [9 empty] | wt.%
|  |  | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | [9 empty] | 3
3
| 1 | 1 | 0.47 | 0.45 | 0.46 | 0.45623 | 0.40522 | [9 empty] | 1
| 1 | 2 | 0.45 | 0.45 | 0.43 | 0.45033 | 0.40852 | [9 empty] | 0
| 1 | 3 | 0.48 | 0.44 | 0.43 | 0.45976 | 0.40043 | [9 empty] | 16
| 1 | 4 | 0.46 | 0.44 | 0.43 | 0.45863 | 0.40431 | [9 empty] | 0.44094
| 1 | 5 | 0.47 | 0.44 | 0.43 | 0.46204 | 0.41085 | [9 empty] | 132
| 1 | 6 | 0.46 | 0.45 | 0.43 | 0.46463 | 0.40725
| Mean |  | 0.465 | 0.445 | 0.435 | 0.4586 | 0.4061
| Median |  | 0.465 | 0.445 | 0.43 | 0.45919 | 0.40623
| Std Dev. |  | 0.010488 | 0.0054772 | 0.012247 | 0.00497 | 0.0036289
| Rel.Std.Dev. |  | 2.256% | 1.231% | 2.816% | 1.084% | 0.8936%
| PDM3 |  | 5.457% | 0.9208% | -1.347% | 4.006% | -7.902%
| Z-Score (Absolute) |  | 0.99 | 0 | 0.5 | 0.67 | 1.93
| Analytical results for TOT_XRF in OREAS 931b (Indicative Value 99.92 wt.%) | [15 empty] | Indicative
TOT_XRF | Round | Replicate | Lab | [13 empty] | 1
| No | No | 03 | [13 empty] | wt.%
|  |  | BF*XRF | [13 empty] | 2
2
| 1 | 1 | 99.83 | [13 empty] | 1
| 1 | 2 | 100.06 | [13 empty] | 12
| 1 | 3 | 99.75 | [13 empty] | 16
| 1 | 4 | 99.71 | [13 empty] | 99.9167
| 1 | 5 | 100.06 | [13 empty] | 18
| 1 | 6 | 100.09
| Mean |  | 99.9167
| Median |  | 99.945
| Std Dev. |  | 0.1727
| Rel.Std.Dev. |  | 0.1728%
| PDM3 |  | -3.331e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 931b (Indicative Value < 42 ppm) | [15 empty] | Indicative
U | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 1
1
| 1 | 1 | < 42 | [13 empty] | 1
| 1 | 2 | < 42 | [13 empty] | 13
| 1 | 3 | < 42 | [13 empty] | 16
| 1 | 4 | < 42 | [13 empty] | < 42
| 1 | 5 | < 42 | [13 empty] | 19
| 1 | 6 | < 42
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for V2O5 in OREAS 931b (Indicative Value 282 ppm) | [15 empty] | Indicative
V2O5 | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 19 | 24 | [12 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 0
0
| 1 | 1 | 306.6 | 251.375 | [12 empty] | 1
| 1 | 2 | 381.05 | 242.34 | [12 empty] | 14
| 1 | 3 | 307.45 | 183.402 | [12 empty] | 16
| 1 | 4 | 328.5 | 207.34 | [12 empty] | 281.8598
| 1 | 5 | 366.75 | 249.3 | [12 empty] | 20
| 1 | 6 | 346.65 | 211.56
| Mean |  | 339.5 | 224.2195
| Median |  | 337.575 | 226.95
| Std Dev. |  | 30.8243 | 27.59
| Rel.Std.Dev. |  | 9.079% | 12.3%
| PDM3 |  | 20.45% | -20.45%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for W in OREAS 931b (Indicative Value < 80 ppm) | [15 empty] | Indicative
W | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 0
0
| 1 | 1 | < 80 | [13 empty] | 1
| 1 | 2 | < 80 | [13 empty] | 15
| 1 | 3 | < 80 | [13 empty] | 16
| 1 | 4 | < 80 | [13 empty] | < 80
| 1 | 5 | < 80 | [13 empty] | 21
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 931b (Indicative Value < 39 ppm) | [15 empty] | Indicative
Y | Round | Replicate | Lab | [13 empty] | 1
| No | No | 24 | [13 empty] | ppm
|  |  | OxBF*XRF | [13 empty] | 1
1
| 1 | 1 | 40.3815 | [13 empty] | 1
| 1 | 2 | < 39 | [13 empty] | 16
| 1 | 3 | < 39 | [13 empty] | 16
| 1 | 4 | < 39 | [13 empty] | < 39
| 1 | 5 | < 39 | [13 empty] | 22
| 1 | 6 | < 39
| Mean |  | 40.3815
| Median |  | 40.3815
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 931b (Certified Value 290 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 14 | 16 | 19 | 24 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | OxBF*XRF | BF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | BF*XRF | OxBF*XRF | 0
0
| 1 | 1 | 190 | 300 | 490 | 280 | 300 | 260 | 290 | 100 | 320 | 3200 | 320 | 300 | 74.4 | 252.829 | 1
| 1 | 2 | 190 | 300 | 510 | 270 | 280 | 260 | 290 | 200 | 320 | 2700 | 320 | 300 | 85.1 | 247.5241 | 0
| 1 | 3 | 200 | 280 | 510 | 280 | 320 | 260 | 280 | 100 | 330 | 100 | 320 | 300 | 74.7 | 252.1057 | 16
| 1 | 4 | 190 | 300 | 530 | 280 | 290 | 260 | 290 | 200 | 320 | 1000 | 320 | 300 | 75.1 | 261.5069 | 289.7671
| 1 | 5 | 210 | 300 | 530 | 290 | 300 | 260 | 290 | 200 | 320 | 300 | 240 | 300 | 86.7 | 245.9354 | 133
| 1 | 6 | 200 | 310 | 590 | 280 | 300 | 250 | 300 | 200 | 330 | < 100 | 320 | 300 | 92.7 | 249.8339
| Mean |  | 196.6667 | 298.3333 | 526.6667 | 280 | 298.3333 | 258.3333 | 290 | 166.6667 | 323.3333 | 1460 | 306.6667 | 300 | 81.45 | 251.6225
| Median |  | 195 | 300 | 520 | 280 | 300 | 260 | 290 | 200 | 320 | 1000 | 320 | 300 | 80.1 | 250.9698
| Std Dev. |  | 8.165 | 9.8319 | 34.448 | 6.3246 | 13.2916 | 4.0825 | 6.3246 | 51.6398 | 5.164 | 1411.7365 | 32.6599 | 0 | 7.785 | 5.5087
| Rel.Std.Dev. |  | 4.152% | 3.296% | 6.541% | 2.259% | 4.455% | 1.58% | 2.181% | 30.98% | 1.597% | 96.69% | 10.65% | 0% | 9.558% | 2.189%
| PDM3 |  | -32.13% | 2.956% | 81.76% | -3.371% | 2.956% | -10.85% | 0.08039% | -42.48% | 11.58% | 403.9% | 5.832% | 3.531% | -71.89% | -13.16%
| Z-Score (Absolute) |  | 3.03 | 0.13 | 7.24 | 0.44 | 0.13 | 1.12 | 0.13 | NA | 0.91 | NA | 0.43 | 0.18 | 6.62 | 1.32
Results from laboratories 10 and 12 were removed due to their 100 ppm reading resolution.
| Analytical results for ZrO2 in OREAS 931b (Indicative Value 45 ppm) | [15 empty] | Indicative
ZrO2 | Round | Replicate | Lab | Lab | [12 empty] | 1
| No | No | 19 | 24 | [12 empty] | ppm
|  |  | BF*XRF | OxBF*XRF | [12 empty] | 1
1
| 1 | 1 | < 50 | < 500 | [12 empty] | 1
| 1 | 2 | 58.65 | < 500 | [12 empty] | 17
| 1 | 3 | < 50 | < 500 | [12 empty] | 16
| 1 | 4 | 64.2 | < 500 | [12 empty] | 45.0333
| 1 | 5 | 44.3 | < 500 | [12 empty] | 23
| 1 | 6 | 53.05 | < 500
| Mean |  | 55.05
| Median |  | 55.85
| Std Dev. |  | 8.4901
| Rel.Std.Dev. |  | 15.42%
| PDM3 |  | 22.24%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C774: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 931b (Indicative Value 17.2 ppm) | [22 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [18 empty] | 1
| No | No | 14 | 15 | 20 | [18 empty] | ppm
|  |  | PF*OES/MS | PF*MS | PF*MS | [18 empty] | 1
1
| 1 | 1 | 18 | 18 | 14 | [18 empty] | 1
| 1 | 2 | 17 | 14 | 15 | [18 empty] | 19
| 1 | 3 | 18 | < 10 | 16 | [18 empty] | 16
| 1 | 4 | 18 | 17 | 16 | [18 empty] | 17.1778
| 1 | 5 | 18 | 17 | 16 | [18 empty] | 25
| 1 | 6 | 18 | 25 | 16
| Mean |  | 17.8333 | 18.2 | 15.5
| Median |  | 18 | 17 | 16
| Std Dev. |  | 0.40825 | 4.0866 | 0.83666
| Rel.Std.Dev. |  | 2.289% | 22.45% | 5.398%
| PDM3 |  | 3.816% | 5.951% | -9.767%
| Z-Score (Absolute) |  | 2.47 | 0 | 0.67
| Analytical results for Al in OREAS 931b (Certified Value 5.73 wt.%) | [22 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 2
3
| 1 | 1 | 5.82 | 5.7945 | 5.56 | 5.81 | 5.689 | 5.663 | 5.875 | 5.822 | 5.742 | 5.35 | 5.62 | 6.017 | 5.6 | 5.32 | 5.74 | 5.78 | 5.6 | 5.7587 | 5.62 | 6.298 | 5.6592 | 1
| 1 | 2 | 5.84 | 5.9412 | 5.68 | 5.83 | 5.716 | 5.716 | 5.769 | 5.848 | 5.716 | 5.74 | 5.59 | 5.531 | 5.62 | 5.3 | 5.68 | 5.84 | 5.74 | 5.7867 | 5.51 | 6.033 | 5.8845 | 0
| 1 | 3 | 5.84 | 5.7782 | 5.61 | 5.95 | 5.689 | 5.716 | 5.795 | 5.848 | 5.663 | 5.46 | 5.53 | 5.779 | 5.6 | 5.42 | 5.76 | 5.83 | 5.82 | 5.7587 | 5.63 | 6.086 | 5.7028 | 16
| 1 | 4 | 5.86 | 5.8875 | 5.62 | 5.83 | 5.663 | 5.663 | 5.557 | 5.848 | 5.716 | 5.41 | 5.68 | 5.853 | 5.53 | 5.52 | 5.69 | 5.69 | 5.81 | 5.7587 | 5.68 | 5.822 | 5.9363 | 5.7319
| 1 | 5 | 5.89 | 5.8145 | 5.6 | 5.83 | 5.742 | 5.689 | 5.822 | 5.954 | 5.769 | 5.5 | 5.67 | 5.996 | 5.63 | 5.37 | 5.72 | 5.79 | 5.83 | 5.9931 | 5.63 | 6.192 | 5.9006 | 135
| 1 | 6 | 5.89 | 5.9551 | 5.72 | 5.56 | 5.689 | 5.689 | 5.927 | 5.927 | 5.716 | 5.54 | 5.55 | 5.875 | 5.54 | 5.39 | 5.68 | 5.75 | 5.78 | 5.574 | 5.63 | 5.927 | 5.8995
| Mean |  | 5.8567 | 5.8618 | 5.6317 | 5.8017 | 5.698 | 5.6893 | 5.7908 | 5.8745 | 5.7203 | 5.5 | 5.6067 | 5.8418 | 5.5867 | 5.3867 | 5.7117 | 5.78 | 5.7633 | 5.7717 | 5.6167 | 6.0597 | 5.8305
| Median |  | 5.85 | 5.851 | 5.615 | 5.83 | 5.689 | 5.689 | 5.8085 | 5.848 | 5.716 | 5.48 | 5.605 | 5.864 | 5.6 | 5.38 | 5.705 | 5.785 | 5.795 | 5.7587 | 5.63 | 6.0595 | 5.892
| Std Dev. |  | 0.028752 | 0.076714 | 0.058109 | 0.12875 | 0.027306 | 0.023704 | 0.12788 | 0.052801 | 0.035161 | 0.1352 | 0.061536 | 0.17677 | 0.041793 | 0.078909 | 0.033714 | 0.055136 | 0.08641 | 0.13329 | 0.056451 | 0.17309 | 0.11786
| Rel.Std.Dev. |  | 0.4909% | 1.309% | 1.032% | 2.219% | 0.4792% | 0.4166% | 2.208% | 0.8988% | 0.6147% | 2.458% | 1.098% | 3.026% | 0.7481% | 1.465% | 0.5903% | 0.9539% | 1.499% | 2.309% | 1.005% | 2.856% | 2.021%
| PDM3 |  | 2.177% | 2.267% | -1.748% | 1.218% | -0.5911% | -0.7423% | 1.029% | 2.488% | -0.2015% | -4.045% | -2.185% | 1.918% | -2.533% | -6.023% | -0.3527% | 0.8395% | 0.5487% | 0.6938% | -2.01% | 5.719% | 1.721%
| Z-Score (Absolute) |  | 0.8 | 0.85 | 1.13 | 0.33 | 0.56 | 0.63 | 0.24 | 0.95 | 0.37 | 2.26 | 1.34 | 0.67 | 1.52 | 3.23 | 0.44 | 0.14 | 0 | 0.07 | 1.26 | 2.54 | 0.58
| Analytical results for As in OREAS 931b (Certified Value 97 ppm) | [22 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 |  |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES |  |  | 0
1
| 1 | 1 | 104 | 101.9438 | 72 | 110 | 100 | 100 | 200 | 100 | 200 | 58 | 77 | 89 | 103 | 100 | 90 | < 100 | 107.449 | 94 | < 100 |  |  | 1
| 1 | 2 | 107 | 101.8529 | 79 | 100 | 100 | 100 | 100 | 200 | 200 | 82 | 62 | 87 | 102 | 200 | 90 | < 100 | 100.8224 | 93 | < 100 |  |  | 0
| 1 | 3 | 104 | 101.0936 | 69 | 120 | 100 | < 100 | 300 | 200 | 200 | 86 | 58 | 107 | 104 | 200 | 98 | < 100 | 98.5076 | 92 | < 100 |  |  | 16
| 1 | 4 | 103 | 101.9702 | 85 | 100 | 200 | 100 | 200 | 200 | 200 | 81 | 50 | 96 | 97 | < 100 | 95 | < 100 | 106.8691 | 94 | 100 |  |  | 96.9533
| 1 | 5 | 103 | 104.8327 | 90 | 110 | 100 | 100 | 200 | 200 | 200 | 89 | 38 | 100 | 102 | 100 | 90 | < 100 | 93.3924 | 92 | < 100 |  |  | 136
| 1 | 6 | 103 | 104.815 | 84 | 110 | 100 | 100 | < 100 | 200 | 200 | 75 | 52 | 99 | 102 | 200 | 79 | < 100 | 103.2047 | 90 | 100
| Mean |  | 104 | 102.7514 | 79.8333 | 108.3333 | 116.6667 | 100 | 200 | 183.3333 | 200 | 78.5 | 56.1667 | 96.3333 | 101.6667 | 160 | 90.3333 |  | 101.7076 | 92.5 | 100
| Median |  | 103.5 | 101.957 | 81.5 | 110 | 100 | 100 | 200 | 200 | 200 | 81.5 | 55 | 97.5 | 102 | 200 | 90 |  | 102.0135 | 92.5 | 100
| Std Dev. |  | 1.5492 | 1.6376 | 8.085 | 7.5277 | 40.8248 | 0 | 70.7107 | 40.8248 | 0 | 11.1131 | 13.0907 | 7.4207 | 2.4221 | 54.7723 | 6.4704 |  | 5.3294 | 1.5166 | 0
| Rel.Std.Dev. |  | 1.49% | 1.594% | 10.13% | 6.949% | 34.99% | 0% | 35.36% | 22.27% | 0% | 14.16% | 23.31% | 7.703% | 2.382% | 34.23% | 7.163% |  | 5.24% | 1.64% | 0%
| PDM3 |  | 7.268% | 5.98% | -17.66% | 11.74% | 20.33% | 3.142% | 106.3% | 89.09% | 106.3% | -19.03% | -42.07% | -0.6394% | 4.862% | 65.03% | -6.828% |  | 4.904% | -4.593% | 3.142%
| Z-Score (Absolute) |  | 0.1 | 0.05 | 0.98 | 0.3 | 0.67 | 0.45 | 3.3 | 3.67 | 4.42 | 1.04 | 2.05 | 0.24 | 0 | 1.8 | 0.51 | 2.32 | 0 | 0.41 | 1.57
| Analytical results for B in OREAS 931b (Certified Value 114 ppm) | [22 empty] | Certified
B | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 10; 04: 14; 05: 15; 06: 17; 07: 24 |  |  | ppm
|  |  | 01: PF*OES; 02: PF*OES; 03: PF*OES; 04: PF*OES/MS; 05: PF*OES; 06: PF*OES/MS; 07: PF*MS |  |  | 0
0
| 1 | 1 | 01: 106.2556; 02: 132; 03: < 1000; 04: 49; 05: 103; 06: 110; 07: 148.8861 |  |  | 1
| 1 | 2 | 01: 117.265; 02: 127; 03: < 1000; 04: 40; 05: 93; 06: 110; 07: 108.5403 |  |  | 1
| 1 | 3 | 01: 116.356; 02: 127; 03: < 1000; 04: 59; 05: 103; 06: 110; 07: 118.9446 |  |  | 16
| 1 | 4 | 01: 107.0445; 02: 128; 03: < 1000; 04: 59; 05: 99; 06: 110; 07: 119.3461 |  |  | 113.7586
| 1 | 5 | 01: 111.5177; 02: 132; 03: < 1000; 04: 59; 05: 103; 06: 110; 07: 115.4176 |  |  | 137
| 1 | 6 | 01: 117.387; 02: 130; 03: < 1000; 04: 55; 05: 104; 06: 110; 07: 117.6936
| Mean |  | 01: 112.6376; 02: 129.3333; 04: 53.5; 05: 100.8333; 06: 110; 07: 121.4714
| Median |  | 01: 113.9368; 02: 129; 04: 57; 05: 103; 06: 110; 07: 118.3191
| Std Dev. |  | 01: 5.1198; 02: 2.3381; 04: 7.6877; 05: 4.2151; 06: 0; 07: 14.0043
| Rel.Std.Dev. |  | 01: 4.545%; 02: 1.808%; 04: 14.37%; 05: 4.18%; 06: 0%; 07: 11.53%
| PDM3 |  | 01: -0.9854%; 02: 13.69%; 04: -52.97%; 05: -11.36%; 06: -3.304%; 07: 6.78%
| Z-Score (Absolute) |  | 01: 0; 02: 0.95; 03: 22.13; 04: 3.38; 05: 0.67; 06: 0.15; 07: 0.5
| Analytical results for Ba in OREAS 931b (Certified Value 1739 ppm) | [22 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [7 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | [7 empty] | 0
0
| 1 | 1 | 1821 | 1755.8933 | 1764 | 1680 | 1663 | 1748 | 1730 | 1670 | 1737 | 1680 | 1748 | 1930 | 1750 | 1856.0239 | [7 empty] | 1
| 1 | 2 | 1805 | 1782.7813 | 1747 | 1690 | 1806 | 1728 | 1691 | 1660 | 1712 | 1690 | 1751 | 1820 | 1699 | 1899.7237 | [7 empty] | 3
| 1 | 3 | 1825 | 1698.2067 | 1782 | 1760 | 1750 | 1683 | 1718 | 1680 | 1761 | 1740 | 1791 | 1830 | 1738 | 1845.7029 | [7 empty] | 16
| 1 | 4 | 1802 | 1774.3038 | 1745 | 1650 | 1754 | 1731 | 1721 | 1640 | 1722 | 1740 | 1710 | 1810 | 1748 | 1890.5404 | [7 empty] | 1739.3893
| 1 | 5 | 1855 | 1770.9747 | 1761 | 1670 | 1754 | 1672 | 1627 | 1680 | 1727 | 1730 | 1797 | 1890 | 1734 | 1904.3358 | [7 empty] | 138
| 1 | 6 | 1840 | 1698.2064 | 1740 | 1700 | 1761 | 1715 | 1659 | 1650 | 1734 | 1660 | 1785 | 1770 | 1734 | 1913.5274
| Mean |  | 1824.6667 | 1746.7277 | 1756.5 | 1691.6667 | 1748 | 1712.8333 | 1691 | 1663.3333 | 1732.1667 | 1706.6667 | 1763.6667 | 1841.6667 | 1733.8333 | 1884.9757
| Median |  | 1823 | 1763.434 | 1754 | 1685 | 1754 | 1721.5 | 1704.5 | 1665 | 1730.5 | 1710 | 1768 | 1825 | 1736 | 1895.1321
| Std Dev. |  | 20.3437 | 38.5768 | 15.6301 | 37.6386 | 46.5489 | 29.5257 | 40.6694 | 16.3299 | 16.7023 | 34.448 | 33.4644 | 58.1091 | 18.4002 | 27.6371
| Rel.Std.Dev. |  | 1.115% | 2.209% | 0.8898% | 2.225% | 2.663% | 1.724% | 2.405% | 0.9818% | 0.9642% | 2.018% | 1.897% | 3.155% | 1.061% | 1.466%
| PDM3 |  | 4.903% | 0.4219% | 0.9837% | -2.744% | 0.495% | -1.527% | -2.782% | -4.373% | -0.4152% | -1.881% | 1.396% | 5.88% | -0.3194% | 8.37%
| Z-Score (Absolute) |  | 1.86 | 0.14 | 0.36 | 1.07 | 0.17 | 0.61 | 1.09 | 1.7 | 0.18 | 0.74 | 0.52 | 2.24 | 0.14 | 3.2
| Analytical results for Be in OREAS 931b (Certified Value 2.06 ppm) | [22 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 10 | 12 | 13 | 14 | 17 | 19 | 20 | 22 | 24 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | [9 empty] | 2
3
| 1 | 1 | 2.3194 | 2 | 2 | < 5 | 2 | 2 | < 5 | < 3 | 1.6806 | < 5 | 2.1 | 1.8829 | [9 empty] | 1
| 1 | 2 | 2.2678 | 2 | 2 | < 5 | 2 | 2 | < 5 | < 3 | 1.8729 | < 5 | 2.2 | 2.6864 | [9 empty] | 0
| 1 | 3 | 2.0764 | 2 | 2 | < 5 | 2 | 2 | < 5 | < 3 | 1.9415 | < 5 | 2.1 | 2.1831 | [9 empty] | 16
| 1 | 4 | 2.3855 | 3 | 2 | < 5 | 2 | 2 | < 5 | < 3 | 2.3279 | < 5 | 2 | 2.558 | [9 empty] | 2.0569
| 1 | 5 | 2.3653 | 2 | 2 | < 5 | 2 | 2 | < 5 | < 3 | 2.0191 | < 5 | 2.2 | 1.7184 | [9 empty] | 139
| 1 | 6 | 2.0047 | 2 | 2 | < 5 | 2 | 2 | < 5 | < 3 | 2.2183 | < 5 | 2.1 | 1.5243
| Mean |  | 01: 2.2365; 02: 2.1667; 03: 2; 05: 2; 06: 2; 09: 2.01; 12: 2.1167; 14: 2.0922
| Median |  | 01: 2.2936; 02: 2; 03: 2; 05: 2; 06: 2; 09: 1.9803; 12: 2.1; 14: 2.033
| Std Dev. |  | 01: 0.15876; 02: 0.40825; 03: 0; 05: 0; 06: 0; 09: 0.23516; 12: 0.075277; 14: 0.46571
| Rel.Std.Dev. |  | 01: 7.098%; 02: 18.84%; 03: 0%; 05: 0%; 06: 0%; 09: 11.7%; 12: 3.556%; 14: 22.26%
| PDM3 |  | 01: 8.731%; 02: 5.335%; 03: -2.768%; 05: -2.768%; 06: -2.768%; 09: -2.28%; 12: 2.904%; 14: 1.715%
| Z-Score (Absolute) |  | 0.85 | 0.4 | 0.67 | 2.55 | 0.67 | 0.67 | 2.55 | 3.9 | 0.61 | 2.55 | 0.08 | 0.08
| Analytical results for Bi in OREAS 931b (Certified Value 216 ppm) | [22 empty] | Certified
Bi | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 12; 05: 14; 06: 15; 07: 17; 08: 19; 09: 20; 10: 24 |  |  | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 08: PF*MS; 09: PF*MS; 10: PF*MS |  |  | 0
0
| 1 | 1 | 01: 214.2543; 02: 215; 03: 219; 04: 27; 05: 221; 06: 228.8; 07: 209; 08: 122.0648; 09: 224; 10: 195.3748 |  |  | 1
| 1 | 2 | 01: 214.1131; 02: 217.1; 03: 235; 04: 36.1; 05: 217; 06: 225.2; 07: 194; 08: 115.0083; 09: 213; 10: 214.3962 |  |  | 8
| 1 | 3 | 01: 211.9819; 02: 214.3; 03: 229; 04: 42.6; 05: 216; 06: 234.7; 07: 195; 08: 148.4747; 09: 218; 10: 204.2443 |  |  | 16
| 1 | 4 | 01: 214.8559; 02: 218.2; 03: 231; 04: 48.3; 05: 215; 06: 231; 07: 197; 08: 137.2005; 09: 220; 10: 194.8894 |  |  | 215.9555
| 1 | 5 | 01: 214.4717; 02: 220.4; 03: 236; 04: 44.7; 05: 219; 06: 235.4; 07: 196; 08: 116.081; 09: 217; 10: 217.9859 |  |  | 140
| 1 | 6 | 01: 211.8936; 02: 216.8; 03: 224; 04: 72.2; 05: 211; 06: 228.1; 07: 193; 08: 137.7462; 09: 216; 10: 221.4048
| Mean |  | 01: 213.5951; 02: 216.9667; 03: 229; 04: 45.15; 05: 216.5; 06: 230.5333; 07: 197.3333; 08: 129.4293; 09: 218; 10: 208.0492
| Median |  | 01: 214.1837; 02: 216.95; 03: 230; 04: 43.65; 05: 216.5; 06: 229.9; 07: 195.5; 08: 129.6327; 09: 217.5; 10: 209.3202
| Std Dev. |  | 01: 1.3083; 02: 2.2061; 03: 6.5422; 04: 15.2293; 05: 3.4496; 06: 3.9657; 07: 5.8878; 08: 13.6577; 09: 3.7417; 10: 11.5386
| Rel.Std.Dev. |  | 01: 0.6125%; 02: 1.017%; 03: 2.857%; 04: 33.73%; 05: 1.593%; 06: 1.72%; 07: 2.984%; 08: 10.55%; 09: 1.716%; 10: 5.546%
| PDM3 |  | 01: -1.093%; 02: 0.4682%; 03: 6.04%; 04: -79.09%; 05: 0.2521%; 06: 6.75%; 07: -8.623%; 08: -40.07%; 09: 0.9467%; 10: -3.661%
| Z-Score (Absolute) |  | 01: 0.09; 02: 0.12; 03: 0.9; 04: 10.94; 05: 0.09; 06: 1; 07: 1.14; 08: 5.51; 09: 0.19; 10: 0.45
| Analytical results for Ca in OREAS 931b (Certified Value 0.096 wt.%) | [22 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS |  | 3
3
| 1 | 1 | 0.09 | < 0.1 | < 0.1 | 0.25 | 0.071 | < 0.05 | 0.157 | 0.064 | 0.15 | < 0.1 | 0.1 | 0.121 | 0.16 | 0.3 | 0.12 | 0.08 | 0.09 | < 0.36 | 0.1 | 0.11235 |  | 1
| 1 | 2 | 0.09 | < 0.1 | 0.1 | 0.26 | 0.071 | < 0.05 | 0.121 | 0.143 | 0.079 | < 0.1 | 0.2 | 0.121 | 0.14 | 0.2 | 0.08 | 0.07 | 0.09 | < 0.36 | < 0.1 | 0.10944 |  | 0
| 1 | 3 | 0.09 | < 0.1 | < 0.1 | 0.25 | 0.057 | < 0.05 | 0.121 | 0.064 | 0.2 | < 0.1 | 0.1 | 0.114 | 0.1 | 0.2 | 0.11 | 0.09 | 0.1 | < 0.36 | 0.1 | 0.10936 |  | 16
| 1 | 4 | 0.07 | < 0.1 | 0.1 | 0.25 | 0.057 | < 0.05 | 0.093 | 0.086 | 0.079 | < 0.1 | 0.2 | 0.121 | 0.08 | 0.2 | 0.11 | 0.05 | 0.09 | < 0.36 | 0.1 | 0.10755 |  | 0.096169
| 1 | 5 | 0.1 | < 0.1 | 0.1 | 0.25 | 0.071 | < 0.05 | 0.129 | 0.086 | 0.057 | < 0.1 | 0.2 | 0.121 | 0.08 | 0.2 | 0.08 | 0.09 | 0.12 | < 0.36 | 0.1 | 0.11104 |  | 141
| 1 | 6 | 0.1 | < 0.1 | 0.1 | 0.25 | 0.079 | < 0.05 | 0.229 | 0.079 | 0.071 | < 0.1 | 0.1 | 0.121 | 0.08 | 0.3 | 0.1 | 0.08 | 0.1 | < 0.36 | < 0.1 | 0.11434
| Mean |  | 0.09 |  | 0.1 | 0.25167 | 0.067667 |  | 0.14167 | 0.087 | 0.106 |  | 0.15 | 0.11983 | 0.10667 | 0.23333 | 0.1 | 0.076667 | 0.098333 |  | 0.1 | 0.11068
| Median |  | 0.09 |  | 0.1 | 0.25 | 0.071 |  | 0.125 | 0.0825 | 0.079 |  | 0.15 | 0.121 | 0.09 | 0.2 | 0.105 | 0.08 | 0.095 |  | 0.1 | 0.11024
| Std Dev. |  | 0.010954 |  | 0 | 0.0040825 | 0.0088242 |  | 0.047424 | 0.029189 | 0.056313 |  | 0.054772 | 0.0028577 | 0.035024 | 0.05164 | 0.016733 | 0.015055 | 0.01169 |  | 0 | 0.0024249
| Rel.Std.Dev. |  | 12.17% |  | 0% | 1.622% | 13.04% |  | 33.48% | 33.55% | 53.13% |  | 36.51% | 2.385% | 32.83% | 22.13% | 16.73% | 19.64% | 11.89% |  | 0% | 2.191%
| PDM3 |  | -6.415% |  | 3.984% | 161.7% | -29.64% |  | 47.31% | -9.534% | 10.22% |  | 55.98% | 24.61% | 10.92% | 142.6% | 3.984% | -20.28% | 2.251% |  | 3.984% | 15.09%
| Z-Score (Absolute) |  | 0.28 | 1.52 | NA | 4.71 | 0.97 | 2.29 | 1.32 | 0.38 | 0.21 | 1.52 | NA | 0.65 | 0.23 | NA | 0.03 | 0.7 | 0.03 | 2.46 | NA | 0.36
Results from laboratories 3, 12, 15 and 20 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 931b (Indicative Value 0.64 ppm) | [22 empty] | Indicative
Cd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  | 1
| No | No | 01: 02; 02: 10; 03: 12; 04: 14; 05: 17; 06: 19; 07: 20 |  | ppm
|  |  | 01: PF*MS; 02: PF*OES; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*MS |  | 2
2
| 1 | 1 | 01: < 10; 02: < 10; 03: 0.6; 04: 0.5; 05: < 2; 06: 1.9527; 07: 0.6 |  | 1
| 1 | 2 | 01: < 10; 02: < 10; 03: 0.8; 04: 0.6; 05: < 2; 06: 2.2602; 07: 0.6 |  | 20
| 1 | 3 | 01: < 10; 02: < 10; 03: 0.7; 04: 0.7; 05: < 2; 06: 1.0922; 07: 0.5 |  | 16
| 1 | 4 | 01: < 10; 02: < 10; 03: 0.8; 04: 0.6; 05: < 2; 06: 1.4511; 07: 0.6 |  | 0.63889
| 1 | 5 | 01: < 10; 02: < 10; 03: 0.6; 04: 0.9; 05: < 2; 06: 2.1608; 07: 0.6 |  | 26
| 1 | 6 | 01: < 10; 02: < 10; 03: 0.8; 04: 0.4; 05: < 2; 06: 1.5215; 07: 0.6
| Mean |  | 03: 0.71667; 04: 0.61667; 06: 1.7397; 07: 0.58333
| Median |  | 03: 0.75; 04: 0.6; 06: 1.7371; 07: 0.6
| Std Dev. |  | 03: 0.098319; 04: 0.17224; 06: 0.45688; 07: 0.040825
| Rel.Std.Dev. |  | 03: 13.72%; 04: 27.93%; 06: 26.26%; 07: 6.999%
| PDM3 |  | 03: 12.17%; 04: -3.478%; 06: 172.3%; 07: -8.696%
| Z-Score (Absolute) |  | 01: 6.47; 02: 6.47; 03: 0.46; 04: 0.62; 05: 0; 06: 1.2; 07: 0.67
| Analytical results for Ce in OREAS 931b (Certified Value 76 ppm) | [22 empty] | Certified
Ce | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  | 1
| No | No | 11 | 12 | 13 | 14 | 17 | 18 | 19 | 20 | 24 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [12 empty] | 0
1
| 1 | 1 | 69.599 | 78.7 | 70.9 | 81.2 | 74.1 | 76.8 | 77.7924 | 76.7 | 79.026 | [12 empty] | 1
| 1 | 2 | 71.3833 | 85.5 | 73.7 | 80.7 | 72.1 | 76.2 | 72.7103 | 73.4 | 85.1597 | [12 empty] | 9
| 1 | 3 | 69.4333 | 83.4 | 73 | 78.3 | 72.2 | 77.3 | 74.2856 | 74.8 | 81.7886 | [12 empty] | 16
| 1 | 4 | 71.1364 | 82.1 | 72 | 76.3 | 65.7 | 75.8 | 73.2912 | 75.3 | 78.1041 | [12 empty] | 76.2399
| 1 | 5 | 74.8979 | 80.8 | 74.3 | 77.7 | 74.1 | 80.8 | 69.7143 | 74.6 | 83.1819 | [12 empty] | 142
| 1 | 6 | 74.1453 | 81.9 | 75.2 | 76.7 | 72.9 | 78.2 | 75.1112 | 73.9 | 81.5135
| Mean |  | 71.7659 | 82.0667 | 73.1833 | 78.4833 | 71.85 | 77.5167 | 73.8175 | 74.7833 | 81.4623
| Median |  | 71.2599 | 82 | 73.35 | 78 | 72.55 | 77.05 | 73.7884 | 74.7 | 81.651
| Std Dev. |  | 2.2869 | 2.3036 | 1.5639 | 2.0439 | 3.138 | 1.8159 | 2.6845 | 1.1548 | 2.6058
| Rel.Std.Dev. |  | 3.187% | 2.807% | 2.137% | 2.604% | 4.367% | 2.343% | 3.637% | 1.544% | 3.199%
| PDM3 |  | -5.868% | 7.643% | -4.009% | 2.943% | -5.758% | 1.675% | -3.177% | -1.91% | 6.85%
| Z-Score (Absolute) |  | 0.69 | 1.67 | 0.37 | 0.85 | 0.67 | 0.63 | 0.22 | 0 | 1.54
| Analytical results for Co in OREAS 931b (Certified Value 28.4 ppm) | [22 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS |  | 1
2
| 1 | 1 | 33 | 31.1543 | 29 | 30 | 20 | < 20 | < 20 | 20 | < 20 | 27 | 30 | 27 | 27.7 | 29 | 30 | 28.3 | 28 | 20.027 | 28.5 | 25.8658 |  | 1
| 1 | 2 | 33 | 31.041 | 28 | 30 | 30 | < 20 | < 20 | < 20 | 20 | 35 | 25.4 | 24.6 | 27.7 | 29 | 20 | 28.7 | 29 | 20.4672 | 28.1 | 26.8524 |  | 11
| 1 | 3 | 33 | 29.79 | 28 | 31 | 30 | 20 | < 20 | 20 | < 20 | 34 | 28.2 | 27.2 | 30.6 | 24 | 30 | 28.5 | 28 | 21.2683 | 28.2 | 26.4959 |  | 16
| 1 | 4 | 33 | 30.2348 | 27 | 31 | 20 | < 20 | < 20 | 20 | 20 | 32 | 28.3 | 28.2 | 28.9 | 27 | 20 | 29.2 | 28 | 21.9085 | 28.2 | 26.2118 |  | 28.3806
| 1 | 5 | 33 | 30.8465 | 28 | 30 | 20 | 30 | < 20 | 20 | < 20 | 32 | 29.3 | 26.1 | 31.2 | 33 | 20 | 27.3 | 29 | 20.6528 | 27.9 | 26.669 |  | 143
| 1 | 6 | 33 | 30.3055 | 27 | 29 | 30 | 30 | < 20 | 30 | < 20 | 33 | 24.1 | 27.1 | 29 | 26 | 30 | 28.7 | 29 | 31.0039 | 28 | 26.5525
| Mean |  | 33 | 30.562 | 27.8333 | 30.1667 | 25 | 26.6667 |  | 22 | 20 | 32.1667 | 27.55 | 26.7 | 29.1833 | 28 | 25 | 28.45 | 28.5 | 22.5546 | 28.15 | 26.4412
| Median |  | 33 | 30.576 | 28 | 30 | 25 | 30 |  | 20 | 20 | 32.5 | 28.25 | 27.05 | 28.95 | 28 | 25 | 28.6 | 28.5 | 20.9605 | 28.15 | 26.5242
| Std Dev. |  | 0 | 0.53478 | 0.75277 | 0.75277 | 5.4772 | 5.7735 |  | 4.4721 | 0 | 2.7869 | 2.3054 | 1.2264 | 1.4552 | 3.0984 | 5.4772 | 0.63797 | 0.54772 | 4.1911 | 0.20736 | 0.35213
| Rel.Std.Dev. |  | 0% | 1.75% | 2.705% | 2.495% | 21.91% | 21.65% |  | 20.33% | 0% | 8.664% | 8.368% | 4.593% | 4.986% | 11.07% | 21.91% | 2.242% | 1.922% | 18.58% | 0.7366% | 1.332%
| PDM3 |  | 16.28% | 7.686% | -1.928% | 6.293% | -11.91% | -6.039% |  | -22.48% | -29.53% | 13.34% | -2.927% | -5.922% | 2.828% | -1.341% | -11.91% | 0.2445% | 0.4207% | -20.53% | -0.8125% | -6.833%
| Z-Score (Absolute) |  | 2.21 | 1.11 | 0.11 | 0.94 | NA | NA | 8.1 | 3.62 | NA | 1.83 | 0.24 | 0.62 | 0.5 | 0.03 | NA | 0.17 | 0.19 | 2.47 | 0.03 | 0.73
Results from laboratories 5, 6, 9 and 16 were removed due to their 10 ppm reading resolution.
| Analytical results for Cr in OREAS 931b (Certified Value 91 ppm) | [22 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS |  | 0
0
| 1 | 1 | 76 | 85.1574 | 108 | 90 | 140 | 70 | 140 | 270 | 140 | 104 | 95 | 96 | 80 | 89 | < 100 | 100 | 100 | 152.7134 | 80 | 83.1874 |  | 1
| 1 | 2 | 78 | 85.624 | 116 | 80 | 140 | 70 | 140 | 140 | 140 | 95 | 95 | 89 | 80 | 89 | < 100 | 100 | 100 | 112.7526 | 75 | 92.219 |  | 0
| 1 | 3 | 79 | 85.2283 | 112 | 90 | 140 | 70 | 140 | 140 | 70 | 99 | 95 | 89 | 90 | 93 | < 100 | 100 | 100 | 102.2414 | 78 | 79.1408 |  | 16
| 1 | 4 | 76 | 80.7185 | 100 | 90 | 140 | 70 | 140 | 70 | 140 | 96 | 100 | 89 | 90 | 91 | < 100 | 110 | 100 | 100.1099 | 77 | 83.6451 |  | 91.1811
| 1 | 5 | 78 | 81.2301 | 92 | 90 | 140 | 70 | 140 | 70 | 70 | 94 | 109 | 89 | 90 | 96 | < 100 | 110 | 100 | 96.496 | 79 | 77.6456 |  | 144
| 1 | 6 | 80 | 85.3945 | 97 | 80 | 140 | 70 | 140 | 70 | 140 | 114 | 95 | 89 | 90 | 110 | < 100 | 110 | 100 | 112.9426 | 76 | 85.228
| Mean |  | 77.8333 | 83.8921 | 104.1667 | 86.6667 | 140 | 70 | 140 | 126.6667 | 116.6667 | 100.3333 | 98.1667 | 90.1667 | 86.6667 | 94.6667 |  | 105 | 100 | 112.876 | 77.5 | 83.511
| Median |  | 78 | 85.1928 | 104 | 90 | 140 | 70 | 140 | 105 | 140 | 97.5 | 95 | 89 | 90 | 92 |  | 105 | 100 | 107.497 | 77.5 | 83.4162
| Std Dev. |  | 1.6021 | 2.2716 | 9.3041 | 5.164 | 0 | 0 | 0 | 78.1452 | 36.1478 | 7.607 | 5.6716 | 2.8577 | 5.164 | 7.9666 |  | 5.4772 | 0 | 20.6466 | 1.8708 | 5.1459
| Rel.Std.Dev. |  | 2.058% | 2.708% | 8.932% | 5.958% | 0% | 0% | 0% | 61.69% | 30.98% | 7.582% | 5.777% | 3.169% | 5.958% | 8.415% |  | 5.216% | 0% | 18.29% | 2.414% | 6.162%
| PDM3 |  | -14.64% | -7.994% | 14.24% | -4.951% | 53.54% | -23.23% | 53.54% | 38.92% | 27.95% | 10.04% | 7.661% | -1.113% | -4.951% | 3.823% |  | 15.16% | 9.672% | 23.79% | -15% | -8.412%
| Z-Score (Absolute) |  | 0.74 | 0.43 | 0.6 | NA | NA | 1.23 | NA | 1.69 | 1.11 | 0.41 | 0.3 | 0.12 | NA | 0.12 | 2.17 | NA | NA | 1.05 | 0.76 | 0.45
| Analytical results for Cs in OREAS 931b (Certified Value 4.92 ppm) | [22 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [9 empty] | 2
3
| 1 | 1 | 5.2539 | 4.6 | 5.1 | 5.4 | 4.7 | 5.1 | 4.7 | 4.7 | 4.9 | 4.8387 | 5 | 7.7062 | [9 empty] | 1
| 1 | 2 | 5.3256 | 4.7 | 5.1 | 5.2 | 3.7 | 4.9 | 4.8 | 4.5 | 5 | 4.5245 | 4.8 | 8.4248 | [9 empty] | 0
| 1 | 3 | 5.0609 | 4.8 | 5.1 | 5.2 | 4.9 | 5 | 5.2 | 4.1 | 5 | 4.9948 | 4.9 | 8.0566 | [9 empty] | 16
| 1 | 4 | 4.948 | 4.8 | 4.8 | 5.1 | 4.8 | 4.4 | 4.9 | 5.6 | 4.9 | 4.8854 | 4.9 | 6.0176 | [9 empty] | 4.9188
| 1 | 5 | 4.9698 | 4.9 | 5 | 5.1 | 4.5 | 4.9 | 5 | 5 | 4.9 | 4.5289 | 4.9 | 7.3876 | [9 empty] | 145
| 1 | 6 | 5.0813 | 4.8 | 5.1 | 5.7 | 4.6 | 4.6 | 5 | 4.5 | 4.9 | 4.8661 | 4.9 | 7.9894
| Mean |  | 5.1066 | 4.7667 | 5.0333 | 5.2833 | 4.5333 | 4.8167 | 4.9333 | 4.7333 | 4.9333 | 4.7731 | 4.9 | 7.597
| Median |  | 5.0711 | 4.8 | 5.1 | 5.2 | 4.65 | 4.9 | 4.95 | 4.6 | 4.9 | 4.8524 | 4.9 | 7.8478
| Std Dev. |  | 0.15246 | 0.10328 | 0.12111 | 0.23166 | 0.43205 | 0.26394 | 0.17512 | 0.5164 | 0.05164 | 0.19805 | 0.063246 | 0.84862
| Rel.Std.Dev. |  | 2.985% | 2.167% | 2.406% | 4.385% | 9.531% | 5.48% | 3.55% | 10.91% | 1.047% | 4.149% | 1.291% | 11.17%
| PDM3 |  | 3.818% | -3.092% | 2.329% | 7.412% | -7.836% | -2.076% | 0.2964% | -3.77% | 0.2964% | -2.962% | -0.3813% | 54.45%
| Z-Score (Absolute) |  | 0.87 | 0.69 | 0.54 | 1.68 | 1.76 | 0.46 | 0.08 | 0.84 | 0.08 | 0.66 | 0.08 | 12.31
| Analytical results for Cu in OREAS 931b (Certified Value 3.96 wt.%) | [22 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | 2
3
| 1 | 1 | 4.24 | 3.9451 | 3.9272 | 3.81 | 3.84 | 3.86 | 4.06 | 4.02 | 3.97 | 3.6871 | 4.0125 | 3.96 | 4.0905 | 3.917 | 3.99 | 3.838 | 3.874 | 3.87 | 4.2 | 3.83 | 3.9938 | 1
| 1 | 2 | 4.21 | 4.0951 | 3.9001 | 3.93 | 3.91 | 3.92 | 3.94 | 4.01 | 3.96 | 4.0238 | 3.9965 | 3.98 | 4.0869 | 3.982 | 3.78 | 3.948 | 3.912 | 3.93 | 4.13 | 3.83 | 4.1146 | 0
| 1 | 3 | 4.28 | 3.969 | 3.7726 | 3.88 | 3.9 | 3.99 | 3.99 | 4.01 | 3.91 | 3.8266 | 3.9288 | 3.92 | 4.2099 | 3.943 | 3.65 | 4.009 | 3.987 | 3.9 | 4.03 | 3.86 | 3.9283 | 16
| 1 | 4 | 4.28 | 4.0425 | 3.7949 | 3.92 | 3.85 | 3.89 | 3.85 | 4.04 | 3.91 | 3.7953 | 4.0527 | 3.97 | 4.2582 | 3.882 | 4.04 | 4.002 | 3.947 | 3.93 | 4.08 | 3.92 | 4.1571 | 3.9577
| 1 | 5 | 4.21 | 4.0453 | 3.8704 | 3.97 | 3.89 | 3.95 | 4.05 | 4.07 | 4 | 3.8832 | 4.0587 | 4.04 | 4.1554 | 4.007 | 4.06 | 3.993 | 3.981 | 3.89 | 4.12 | 3.91 | 4.2008 | 146
| 1 | 6 | 4.29 | 4.0219 | 3.8637 | 3.93 | 3.87 | 3.91 | 4.01 | 4.06 | 3.95 | 3.87 | 3.958 | 3.97 | 4.1586 | 3.933 | 4.11 | 3.969 | 3.852 | 3.93 | 4 | 3.99 | 4.1283
| Mean |  | 4.2517 | 4.0198 | 3.8548 | 3.9067 | 3.8767 | 3.92 | 3.9833 | 4.035 | 3.95 | 3.8477 | 4.0012 | 3.9733 | 4.1599 | 3.944 | 3.9383 | 3.9598 | 3.9255 | 3.9083 | 4.0933 | 3.89 | 4.0871
| Median |  | 4.26 | 4.0322 | 3.867 | 3.925 | 3.88 | 3.915 | 4 | 4.03 | 3.955 | 3.8483 | 4.0045 | 3.97 | 4.157 | 3.938 | 4.015 | 3.981 | 3.9295 | 3.915 | 4.1 | 3.885 | 4.1214
| Std Dev. |  | 0.03656 | 0.054742 | 0.059935 | 0.055377 | 0.028048 | 0.045607 | 0.078401 | 0.025884 | 0.035214 | 0.11113 | 0.051426 | 0.038816 | 0.066828 | 0.044962 | 0.18192 | 0.063817 | 0.055788 | 0.025626 | 0.072572 | 0.06229 | 0.10412
| Rel.Std.Dev. |  | 0.8599% | 1.362% | 1.555% | 1.418% | 0.7235% | 1.163% | 1.968% | 0.6415% | 0.8915% | 2.888% | 1.285% | 0.9769% | 1.606% | 1.14% | 4.619% | 1.612% | 1.421% | 0.6557% | 1.773% | 1.601% | 2.548%
| PDM3 |  | 7.427% | 1.568% | -2.6% | -1.29% | -2.048% | -0.9535% | 0.6468% | 1.952% | -0.1955% | -2.781% | 1.098% | 0.3941% | 5.108% | -0.3471% | -0.4903% | 0.05299% | -0.8145% | -1.248% | 3.426% | -1.711% | 3.27%
| Z-Score (Absolute) |  | 3.97 | 0.92 | 1.25 | 0.57 | 0.97 | 0.4 | 0.44 | 1.12 | 0 | 1.35 | 0.67 | 0.31 | 2.76 | 0.08 | 0.15 | 0.13 | 0.32 | 0.55 | 1.89 | 0.79 | 1.81
| Analytical results for Dy in OREAS 931b (Certified Value 4.12 ppm) | [22 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS | 2
3
| 1 | 1 | 01: 4.1941; 02: 3.76; 03: 3.85; 04: 4.13; 05: 4.1; 06: 4.23; 07: 4.26; 08: 4.8924 | 1
| 1 | 2 | 01: 4.2167; 02: 4.35; 03: 3.72; 04: 4.37; 05: 4.3; 06: 4.17; 07: 4.05; 08: 5.3884 | 4
| 1 | 3 | 01: 4.0669; 02: 4; 03: 3.78; 04: 4.34; 05: 4.1; 06: 4.18; 07: 4.2; 08: 5.2007 | 16
| 1 | 4 | 01: 4.1289; 02: 4.1; 03: 3.83; 04: 4.4; 05: 4; 06: 4.07; 07: 4.26; 08: 4.946 | 4.1169
| 1 | 5 | 01: 4.5137; 02: 4.15; 03: 3.92; 04: 4.32; 05: 3.9; 06: 4.39; 07: 4.14; 08: 4.9505 | 147
| 1 | 6 | 01: 4.2449; 02: 3.8; 03: 3.78; 04: 4.34; 05: 4.2; 06: 4.31; 07: 4.09; 08: 5.5135
| Mean |  | 01: 4.2275; 02: 4.0267; 03: 3.8133; 04: 4.3167; 05: 4.1; 06: 4.225; 07: 4.1667; 08: 5.1486
| Median |  | 01: 4.2054; 02: 4.05; 03: 3.805; 04: 4.34; 05: 4.1; 06: 4.205; 07: 4.17; 08: 5.0756
| Std Dev. |  | 01: 0.15425; 02: 0.22286; 03: 0.069186; 04: 0.095638; 05: 0.14142; 06: 0.11274; 07: 0.088015; 08: 0.2605
| Rel.Std.Dev. |  | 01: 3.649%; 02: 5.535%; 03: 1.814%; 04: 2.216%; 05: 3.449%; 06: 2.668%; 07: 2.112%; 08: 5.06%
| PDM3 |  | 01: 2.686%; 02: -2.193%; 03: -7.375%; 04: 4.851%; 05: -0.4117%; 06: 2.625%; 07: 1.208%; 08: 25.06%
| Z-Score (Absolute) |  | 01: 0.2; 02: 1.05; 03: 2.38; 04: 0.75; 05: 0.6; 06: 0.18; 07: 0.18; 08: 5.93
| Analytical results for Er in OREAS 931b (Certified Value 2.36 ppm) | [22 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS | 2
3
| 1 | 1 | 01: 2.2125; 02: 1.91; 03: 2.19; 04: 2.36; 05: 2.3; 06: 2.42; 07: 2.46; 08: 2.9345 | 1
| 1 | 2 | 01: 2.3897; 02: 2.23; 03: 2.25; 04: 2.32; 05: 2.5; 06: 2.41; 07: 2.26; 08: 3.0646 | 5
| 1 | 3 | 01: 2.3252; 02: 2.6; 03: 2.24; 04: 2.6; 05: 2.4; 06: 2.38; 07: 2.32; 08: 3.1087 | 16
| 1 | 4 | 01: 2.2799; 02: 2.33; 03: 2.29; 04: 2.48; 05: 2.2; 06: 2.32; 07: 2.42; 08: 2.7729 | 2.362
| 1 | 5 | 01: 2.5724; 02: 2.19; 03: 2.25; 04: 2.65; 05: 2.7; 06: 2.39; 07: 2.34; 08: 3.0709 | 148
| 1 | 6 | 01: 2.4108; 02: 2.13; 03: 2.28; 04: 2.51; 05: 2.5; 06: 2.41; 07: 2.39; 08: 2.924
| Mean |  | 01: 2.3651; 02: 2.2317; 03: 2.25; 04: 2.4867; 05: 2.4333; 06: 2.3883; 07: 2.365; 08: 2.9793
| Median |  | 01: 2.3575; 02: 2.21; 03: 2.25; 04: 2.495; 05: 2.45; 06: 2.4; 07: 2.365; 08: 2.9996
| Std Dev. |  | 01: 0.12474; 02: 0.22842; 03: 0.035214; 04: 0.12956; 05: 0.17512; 06: 0.03656; 07: 0.072595; 08: 0.12655
| Rel.Std.Dev. |  | 01: 5.274%; 02: 10.24%; 03: 1.565%; 04: 5.21%; 05: 7.197%; 06: 1.531%; 07: 3.07%; 08: 4.248%
| PDM3 |  | 01: 0.1321%; 02: -5.516%; 03: -4.74%; 04: 5.28%; 05: 3.022%; 06: 1.116%; 07: 0.1285%; 08: 26.14%
| Z-Score (Absolute) |  | 01: 0.09; 02: 1.17; 03: 1.03; 04: 0.89; 05: 0.46; 06: 0.09; 07: 0.09; 08: 4.88
| Analytical results for Eu in OREAS 931b (Certified Value 1.26 ppm) | [22 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS | 2
3
| 1 | 1 | 01: 1.4905; 02: 1.23; 03: 1.22; 04: 1.07; 05: 1.4; 06: 1.32; 07: 1.2; 08: < 2 | 1
| 1 | 2 | 01: 1.3829; 02: 1.34; 03: 1.1; 04: 1.03; 05: 1.3; 06: 1.27; 07: 1.1; 08: < 2 | 6
| 1 | 3 | 01: 1.4204; 02: 1.23; 03: 1.2; 04: 1.14; 05: 1.3; 06: 1.34; 07: 1.11; 08: < 2 | 16
| 1 | 4 | 01: 1.3966; 02: 1.3; 03: 1.25; 04: 1.22; 05: 1.2; 06: 1.32; 07: 1.14; 08: < 2 | 1.2561
| 1 | 5 | 01: 1.4019; 02: 1.16; 03: 1.15; 04: 1.24; 05: 1.2; 06: 1.36; 07: 1.14; 08: < 2 | 149
| 1 | 6 | 01: 1.465; 02: 1.36; 03: 1.19; 04: 1.12; 05: 1.5; 06: 1.33; 07: 1.12; 08: < 2
| Mean |  | 01: 1.4262; 02: 1.27; 03: 1.185; 04: 1.1367; 05: 1.3167; 06: 1.3233; 07: 1.135
| Median |  | 01: 1.4111; 02: 1.265; 03: 1.195; 04: 1.13; 05: 1.3; 06: 1.325; 07: 1.13
| Std Dev. |  | 01: 0.042462; 02: 0.07642; 03: 0.053198; 04: 0.082138; 05: 0.1169; 06: 0.030111; 07: 0.035637
| Rel.Std.Dev. |  | 01: 2.977%; 02: 6.017%; 03: 4.489%; 04: 7.226%; 05: 8.879%; 06: 2.275%; 07: 3.14%
| PDM3 |  | 01: 13.54%; 02: 1.104%; 03: -5.662%; 04: -9.51%; 05: 4.82%; 06: 5.35%; 07: -9.643%
| Z-Score (Absolute) |  | 01: 1.46; 02: 0.31; 03: 0.31; 04: 0.67; 05: 0.66; 06: 0.7; 07: 0.68; 08: 1.67
| Analytical results for Fe in OREAS 931b (Certified Value 8.68 wt.%) | [22 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | 2
3
| 1 | 1 | 8.79 | 9.1578 | 8.53 | 8.68 | 8.533 | 8.568 | 8.883 | 8.358 | 8.778 | 7.79 | 8.62 | 8.995 | 9.54 | 8.16 | 8.6 | 8.76 | 8.64 | 8.74 | 8.72 | 8.63 | 8.8809 | 1
| 1 | 2 | 8.74 | 9.2388 | 8.73 | 8.81 | 8.498 | 8.673 | 8.603 | 8.428 | 8.743 | 8.55 | 8.61 | 8.456 | 9.65 | 8.12 | 8.68 | 8.73 | 8.54 | 8.728 | 8.46 | 8.57 | 9.1971 | 0
| 1 | 3 | 8.74 | 9.1288 | 8.58 | 8.83 | 8.603 | 8.708 | 8.673 | 8.568 | 8.533 | 8.19 | 8.35 | 8.54 | 9.31 | 8.31 | 8.67 | 8.76 | 8.7 | 8.721 | 8.65 | 8.76 | 8.8521 | 16
| 1 | 4 | 8.8 | 9.204 | 8.68 | 8.83 | 8.603 | 8.568 | 8.288 | 8.428 | 8.708 | 8.12 | 8.63 | 8.715 | 9.16 | 8.46 | 8.55 | 8.77 | 8.75 | 8.716 | 8.71 | 8.74 | 9.1063 | 8.6761
| 1 | 5 | 8.8 | 9.0807 | 8.64 | 8.82 | 8.603 | 8.603 | 8.813 | 8.673 | 8.813 | 8.33 | 8.28 | 8.932 | 9.49 | 8.24 | 8.75 | 8.79 | 8.9 | 8.823 | 8.64 | 8.8 | 9.0683 | 150
| 1 | 6 | 8.79 | 9.154 | 8.82 | 8.94 | 8.673 | 8.673 | 8.848 | 8.533 | 8.638 | 8.17 | 8.52 | 8.701 | 9.32 | 8.27 | 8.58 | 8.75 | 8.78 | 8.684 | 8.66 | 8.93 | 9.1695
| Mean |  | 8.7767 | 9.1607 | 8.6633 | 8.8183 | 8.5855 | 8.6322 | 8.6847 | 8.498 | 8.7022 | 8.1917 | 8.5017 | 8.7232 | 9.4117 | 8.26 | 8.6383 | 8.76 | 8.7183 | 8.7353 | 8.64 | 8.7383 | 9.0457
| Median |  | 8.79 | 9.1559 | 8.66 | 8.825 | 8.603 | 8.638 | 8.743 | 8.4805 | 8.7255 | 8.18 | 8.565 | 8.708 | 9.405 | 8.255 | 8.635 | 8.76 | 8.725 | 8.7245 | 8.655 | 8.75 | 9.0873
| Std Dev. |  | 0.028752 | 0.055594 | 0.10443 | 0.082805 | 0.061624 | 0.060284 | 0.2221 | 0.11502 | 0.10244 | 0.25079 | 0.15145 | 0.21118 | 0.17971 | 0.1205 | 0.074677 | 0.02 | 0.12336 | 0.046852 | 0.094021 | 0.12734 | 0.14634
| Rel.Std.Dev. |  | 0.3276% | 0.6069% | 1.205% | 0.939% | 0.7178% | 0.6984% | 2.557% | 1.354% | 1.177% | 3.062% | 1.781% | 2.421% | 1.909% | 1.459% | 0.8645% | 0.2283% | 1.415% | 0.5363% | 1.088% | 1.457% | 1.618%
| PDM3 |  | 1.16% | 5.585% | -0.1468% | 1.64% | -1.044% | -0.506% | 0.09913% | -2.052% | 0.3008% | -5.583% | -2.01% | 0.5429% | 8.479% | -4.796% | -0.4349% | 0.9674% | 0.4872% | 0.6831% | -0.4157% | 0.7177% | 4.26%
| Z-Score (Absolute) |  | 0.72 | 4.42 | 0.38 | 1.12 | 1.12 | 0.67 | 0.17 | 1.97 | 0 | 4.92 | 1.93 | 0.2 | 6.84 | 4.26 | 0.62 | 0.56 | 0.15 | 0.32 | 0.6 | 0.35 | 3.31
| Analytical results for Ga in OREAS 931b (Certified Value 18.3 ppm) | [22 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 24 | [7 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*MS | [7 empty] | 1
1
| 1 | 1 | 22 | 18.1303 | 18 | 21.1 | 8 | 17.6 | 16.8 | 17 | 16.3 | 17 | 18.1496 | 16 | 26 | 21.1448 | [7 empty] | 1
| 1 | 2 | 20 | 18.4176 | 19 | 22.5 | 16 | 17.8 | 17.5 | 17 | 16.6 | 17 | 17.9682 | 16 | 24 | 22.2306 | [7 empty] | 14
| 1 | 3 | 20 | 17.8272 | 18 | 24 | 18 | 18.1 | 20 | 18 | 15.2 | 17 | 17.3868 | 16 | 26 | 21.052 | [7 empty] | 16
| 1 | 4 | 20 | 18.5778 | 18 | 23.3 | 19 | 18.1 | 19.6 | 19 | 16.4 | 17 | 17.7832 | 16 | 25 | 18.7567 | [7 empty] | 18.3076
| 1 | 5 | 22 | 18.5345 | 19 | 20.5 | 17 | 17.6 | 18.8 | 17 | 18.9 | 18 | 17.7935 | 16 | 25 | 21.1791 | [7 empty] | 151
| 1 | 6 | 21 | 17.9978 | 17 | 20.8 | 12 | 18.8 | 17.4 | 18 | 16.2 | 17 | 17.7597 | 16 | 26 | 20.3306
| Mean |  | 20.8333 | 18.2475 | 18.1667 | 22.0333 | 15 | 18 | 18.35 | 17.6667 | 16.6 | 17.1667 | 17.8068 | 16 | 25.3333 | 20.7823
| Median |  | 20.5 | 18.2739 | 18 | 21.8 | 16.5 | 17.95 | 18.15 | 17.5 | 16.35 | 17 | 17.7884 | 16 | 25.5 | 21.0984
| Std Dev. |  | 0.98319 | 0.30762 | 0.75277 | 1.4445 | 4.1952 | 0.45166 | 1.305 | 0.8165 | 1.228 | 0.40825 | 0.25412 | 0 | 0.8165 | 1.1633
| Rel.Std.Dev. |  | 4.719% | 1.686% | 4.144% | 6.556% | 27.97% | 2.509% | 7.112% | 4.622% | 7.398% | 2.378% | 1.427% | 0% | 3.223% | 5.598%
| PDM3 |  | 13.8% | -0.328% | -0.7697% | 20.35% | -18.07% | -1.68% | 0.2317% | -3.501% | -9.327% | -6.232% | -2.735% | -12.6% | 38.38% | 13.52%
| Z-Score (Absolute) |  | 1.55 | 0.09 | 0.05 | 2.22 | 1.73 | 0.05 | 0.15 | 0.23 | 0.83 | 0.52 | 0.16 | 1.17 | 4.07 | 1.52
| Analytical results for Gd in OREAS 931b (Certified Value 4.94 ppm) | [22 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS | 2
3
| 1 | 1 | 01: 5.0888; 02: 4.78; 03: 4.62; 04: 4.9; 05: 4.3; 06: 5; 07: 5.16; 08: 6.1308 | 1
| 1 | 2 | 01: 5.2802; 02: 5.06; 03: 4.51; 04: 4.76; 05: 4.9; 06: 5; 07: 4.77; 08: 6.0921 | 15
| 1 | 3 | 01: 5.1581; 02: 4.67; 03: 4.62; 04: 4.84; 05: 5.5; 06: 5; 07: 4.99; 08: 6.2598 | 16
| 1 | 4 | 01: 5.3981; 02: 4.68; 03: 4.69; 04: 5.29; 05: 4.9; 06: 5; 07: 4.86; 08: 5.4345 | 4.9416
| 1 | 5 | 01: 5.7646; 02: 4.04; 03: 4.56; 04: 5.17; 05: 5.4; 06: 5; 07: 4.84; 08: 6.0019 | 152
| 1 | 6 | 01: 5.4198; 02: 4.6; 03: 4.56; 04: 5.03; 05: 5.3; 06: 5; 07: 4.77; 08: 6.2984
| Mean |  | 01: 5.3516; 02: 4.6383; 03: 4.5933; 04: 4.9983; 05: 5.05; 06: 5; 07: 4.8983; 08: 6.0362
| Median |  | 01: 5.3392; 02: 4.675; 03: 4.59; 04: 4.965; 05: 5.1; 06: 5; 07: 4.85; 08: 6.1114
| Std Dev. |  | 01: 0.24041; 02: 0.33469; 03: 0.06314; 04: 0.20351; 05: 0.44609; 06: 0; 07: 0.15145; 08: 0.31434
| Rel.Std.Dev. |  | 01: 4.492%; 02: 7.216%; 03: 1.375%; 04: 4.072%; 05: 8.834%; 06: 0%; 07: 3.092%; 08: 5.207%
| PDM3 |  | 01: 8.297%; 02: -6.137%; 03: -7.048%; 04: 1.148%; 05: 2.194%; 06: 1.182%; 07: -0.8756%; 08: 22.15%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.69; 03: 0.77; 04: 0; 05: 0.1; 06: NA; 07: 0.19; 08: 1.98
Results from laboratory 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 931b (Certified Value 2.33 ppm) | [22 empty] | Certified
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 12; 05: 14; 06: 17; 07: 19; 08: 20; 09: 24 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*OES/MS; 07: PF*MS; 08: PF*MS; 09: PF*MS | 2
2
| 1 | 1 | 01: 1.2501; 02: 1; 03: 2.4; 04: 2; 05: 2; 06: 2.4; 07: 2.6379; 08: 2; 09: 2.6888 | 1
| 1 | 2 | 01: 1.0077; 02: 1; 03: 2.6; 04: 2; 05: 2; 06: 2.5; 07: 2.3546; 08: 2; 09: 2.9711 | 0
| 1 | 3 | 01: < 1; 02: 2; 03: 2.6; 04: 2; 05: 3; 06: 2.4; 07: 2.4499; 08: 2; 09: 2.6952 | 16
| 1 | 4 | 01: 1.1068; 02: 2; 03: 2.5; 04: 2; 05: 3; 06: 2.3; 07: 2.4813; 08: 2; 09: 2.4577 | 2.3277
| 1 | 5 | 01: 1.1952; 02: 2; 03: 2.8; 04: 2; 05: 3; 06: 2.1; 07: 2.5819; 08: 2; 09: 3.4332 | 153
| 1 | 6 | 01: 1.327; 02: 2; 03: 2.5; 04: 2; 05: 3; 06: 2.1; 07: 2.4601; 08: 2; 09: 3.3177
| Mean |  | 01: 1.1774; 02: 1.6667; 03: 2.5667; 04: 2; 05: 2.6667; 06: 2.3; 07: 2.4943; 08: 2; 09: 2.9273
| Median |  | 01: 1.1952; 02: 2; 03: 2.55; 04: 2; 05: 3; 06: 2.35; 07: 2.4707; 08: 2; 09: 2.8331
| Std Dev. |  | 01: 0.12425; 02: 0.5164; 03: 0.13663; 04: 0; 05: 0.5164; 06: 0.16733; 07: 0.10113; 08: 0; 09: 0.38511
| Rel.Std.Dev. |  | 01: 10.55%; 02: 30.98%; 03: 5.323%; 04: 0%; 05: 19.36%; 06: 7.275%; 07: 4.055%; 08: 0%; 09: 13.16%
| PDM3 |  | 01: -49.42%; 02: -28.4%; 03: 10.27%; 04: -14.08%; 05: 14.56%; 06: -1.19%; 07: 7.157%; 08: -14.08%; 09: 25.76%
| Z-Score (Absolute) |  | 01: 2.78; 02: 1.42; 03: 0.6; 04: 0.67; 05: 0.82; 06: 0; 07: 0.44; 08: 0.67; 09: 1.41
| Analytical results for Hf in OREAS 931b (Indicative Value 2.29 ppm) | [22 empty] | Indicative
Hf | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 13; 02: 14; 03: 17; 04: 20; 05: 24 | ppm
|  |  | 01: BF*OES/MS; 02: PF*OES/MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*MS | 2
2
| 1 | 1 | 01: 2.6; 02: 2; 03: < 10; 04: 2; 05: 2.1532 | 1
| 1 | 2 | 01: 2.8; 02: 2; 03: 10; 04: 2; 05: 2.1486 | 16
| 1 | 3 | 01: 2.6; 02: 3; 03: < 10; 04: 2; 05: 2.1878 | 16
| 1 | 4 | 01: 2.7; 02: 2; 03: < 10; 04: 2; 05: < 2 | 2.2945
| 1 | 5 | 01: 2.7; 02: 3; 03: < 10; 04: 2; 05: 2.6984 | 27
| 1 | 6 | 01: 2.6; 02: 2; 03: < 10; 04: 2; 05: 2.881
| Mean |  | 01: 2.6667; 02: 2.3333; 03: 10; 04: 2; 05: 2.4138
| Median |  | 01: 2.65; 02: 2; 03: 10; 04: 2; 05: 2.1878
| Std Dev. |  | 01: 0.08165; 02: 0.5164; 04: 0; 05: 0.3495
| Rel.Std.Dev. |  | 01: 3.062%; 02: 22.13%; 04: 0%; 05: 14.48%
| PDM3 |  | 01: 16.22%; 02: 1.691%; 03: 335.8%; 04: -12.84%; 05: 5.197%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0; 03: 7.08; 04: 0.67; 05: 0.31
| Analytical results for Ho in OREAS 931b (Certified Value 0.81 ppm) | [22 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS | 2
3
| 1 | 1 | 01: 0.854; 02: 0.74; 03: 0.72; 04: 0.83; 05: 0.8; 06: 0.82; 07: 0.86; 08: < 2 | 1
| 1 | 2 | 01: 0.8177; 02: 0.76; 03: 0.71; 04: 0.81; 05: 0.8; 06: 0.83; 07: 0.83; 08: < 2 | 4
| 1 | 3 | 01: 0.8296; 02: 0.8; 03: 0.7; 04: 0.9; 05: 0.8; 06: 0.8; 07: 0.86; 08: < 2 | 16
| 1 | 4 | 01: 0.8856; 02: 0.77; 03: 0.74; 04: 0.88; 05: 0.8; 06: 0.8; 07: 0.86; 08: < 2 | 0.80914
| 1 | 5 | 01: 0.9035; 02: 0.75; 03: 0.73; 04: 0.86; 05: 0.9; 06: 0.82; 07: 0.82; 08: < 2 | 154
| 1 | 6 | 01: 0.8335; 02: 0.71; 03: 0.76; 04: 0.84; 05: 0.8; 06: 0.8; 07: 0.85; 08: < 2
| Mean |  | 01: 0.85398; 02: 0.755; 03: 0.72667; 04: 0.85333; 05: 0.81667; 06: 0.81167; 07: 0.84667
| Median |  | 01: 0.84375; 02: 0.755; 03: 0.725; 04: 0.85; 05: 0.8; 06: 0.81; 07: 0.855
| Std Dev. |  | 01: 0.034006; 02: 0.030166; 03: 0.021602; 04: 0.033267; 05: 0.040825; 06: 0.013292; 07: 0.017512
| Rel.Std.Dev. |  | 01: 3.982%; 02: 3.996%; 03: 2.973%; 04: 3.898%; 05: 4.999%; 06: 1.638%; 07: 2.068%
| PDM3 |  | 01: 5.542%; 02: -6.691%; 03: -10.19%; 04: 5.462%; 05: 0.9301%; 06: 0.3122%; 07: 4.638%
| Z-Score (Absolute) |  | 01: 0.68; 02: 2.35; 03: 3.22; 04: 0.66; 05: 0.46; 06: 0.61; 07: 0.46; 08: 5.16
| Analytical results for In in OREAS 931b (Certified Value 3.32 ppm) | [22 empty] | Certified
In | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 12; 05: 14; 06: 15; 07: 17; 08: 19; 09: 20; 10: 24 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 08: PF*MS; 09: PF*MS; 10: PF*MS | 2
3
| 1 | 1 | 01: 3.3956; 02: 3.5; 03: 2.98; 04: 3.2; 05: 3.4; 06: 3.4; 07: 3.3; 08: 3.832; 09: 3.3; 10: 4.3982 | 1
| 1 | 2 | 01: 3.3384; 02: 3.4; 03: 3.12; 04: 3.1; 05: 3.4; 06: 3.4; 07: 3.4; 08: 3.8867; 09: 3.2; 10: 4.69 | 0
| 1 | 3 | 01: 3.6028; 02: 3.5; 03: 2.84; 04: 2.9; 05: 3.5; 06: 3.5; 07: 3.1; 08: 4.0154; 09: 3.3; 10: 4.6391 | 16
| 1 | 4 | 01: 3.6608; 02: 3.2; 03: 2.96; 04: 2.8; 05: 3; 06: 3.2; 07: 3.4; 08: 4.0287; 09: 3.3; 10: 4.5187 | 3.3204
| 1 | 5 | 01: 3.6792; 02: 3.4; 03: 3; 04: 2.5; 05: 3.4; 06: 3.3; 07: 3.3; 08: 3.9203; 09: 3.3; 10: 4.7596 | 155
| 1 | 6 | 01: 3.3485; 02: 3.3; 03: 2.89; 04: 2.7; 05: 3.5; 06: 3.4; 07: 3.4; 08: 4.0434; 09: 3.2; 10: 4.4384
| Mean |  | 01: 3.5042; 02: 3.3833; 03: 2.965; 04: 2.8667; 05: 3.3667; 06: 3.3667; 07: 3.3167; 08: 3.9544; 09: 3.2667; 10: 4.574
| Median |  | 01: 3.4992; 02: 3.4; 03: 2.97; 04: 2.85; 05: 3.4; 06: 3.4; 07: 3.35; 08: 3.9678; 09: 3.3; 10: 4.5789
| Std Dev. |  | 01: 0.16024; 02: 0.1169; 03: 0.096695; 04: 0.2582; 05: 0.18619; 06: 0.10328; 07: 0.1169; 08: 0.087064; 09: 0.05164; 10: 0.14456
| Rel.Std.Dev. |  | 01: 4.573%; 02: 3.455%; 03: 3.261%; 04: 9.007%; 05: 5.53%; 06: 3.068%; 07: 3.525%; 08: 2.202%; 09: 1.581%; 10: 3.161%
| PDM3 |  | 01: 5.538%; 02: 1.896%; 03: -10.7%; 04: -13.66%; 05: 1.394%; 06: 1.394%; 07: -0.1114%; 08: 19.1%; 09: -1.617%; 10: 37.76%
| Z-Score (Absolute) |  | 01: 0.78; 02: 0.09; 03: 2.28; 04: 2.84; 05: 0; 06: 0; 07: 0.28; 08: 3.34; 09: 0.57; 10: 6.85
| Analytical results for K in OREAS 931b (Certified Value 2.24 wt.%) | [22 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | 2
3
| 1 | 1 | 2.3 | 2.5369 | 2.25 | 2.08 | 2.192 | 2.167 | 2.258 | 2.258 | 2.266 | 2.18 | 2.3 | 2.125 | 2.29 | 2.18 | 2.22 | 2.3 | 2.1 | 2.2896 | 2.3 | 2.573 | 0.66838 | 1
| 1 | 2 | 2.27 | 2.4487 | 2.26 | 2.15 | 2.25 | 2.192 | 2.217 | 2.275 | 2.258 | 2.27 | 2.3 | 2.1 | 2.25 | 2.16 | 2.24 | 2.4 | 2.2 | 2.3693 | 2.2 | 2.507 | 0.68578 | 0
| 1 | 3 | 2.27 | 2.5682 | 2.26 | 2.13 | 2.225 | 2.241 | 2.183 | 2.258 | 2.291 | 2.2 | 2.3 | 2.15 | 2.32 | 2.19 | 2.17 | 2.3 | 2.2 | 2.3693 | 2.3 | 2.557 | 0.66358 | 16
| 1 | 4 | 2.28 | 2.418 | 2.27 | 2.14 | 2.2 | 2.208 | 2.117 | 2.266 | 2.167 | 2.2 | 2.4 | 2.183 | 2.28 | 2.22 | 2.25 | 2.3 | 2.3 | 2.3277 | 2.3 | 2.474 | 0.68668 | 2.2426
| 1 | 5 | 2.3 | 2.4695 | 2.25 | 2.15 | 2.225 | 2.208 | 2.241 | 2.341 | 2.383 | 2.21 | 2.4 | 2.142 | 2.28 | 2.15 | 2.07 | 2.3 | 2.2 | 2.3634 | 2.3 | 2.59 | 0.68878 | 156
| 1 | 6 | 2.3 | 2.3652 | 2.29 | 2.13 | 2.192 | 2.233 | 2.308 | 2.308 | 2.25 | 2.23 | 2.3 | 2.15 | 2.28 | 2.24 | 2.16 | 2.3 | 2.2 | 2.3319 | 2.3 | 2.532 | 0.69058
| Mean |  | 2.2867 | 2.4678 | 2.2633 | 2.13 | 2.214 | 2.2082 | 2.2207 | 2.2843 | 2.2692 | 2.215 | 2.3333 | 2.1417 | 2.2833 | 2.19 | 2.185 | 2.3167 | 2.2 | 2.3419 | 2.2833 | 2.5388 | 0.68063
| Median |  | 2.29 | 2.4591 | 2.26 | 2.135 | 2.2125 | 2.208 | 2.229 | 2.2705 | 2.262 | 2.205 | 2.3 | 2.146 | 2.28 | 2.185 | 2.195 | 2.3 | 2.2 | 2.3476 | 2.3 | 2.5445 | 0.68623
| Std Dev. |  | 0.015055 | 0.075139 | 0.015055 | 0.026077 | 0.023247 | 0.027022 | 0.065747 | 0.033411 | 0.069821 | 0.031464 | 0.05164 | 0.02779 | 0.022509 | 0.034641 | 0.067157 | 0.040825 | 0.063246 | 0.031625 | 0.040825 | 0.04331 | 0.011569
| Rel.Std.Dev. |  | 0.6584% | 3.045% | 0.6652% | 1.224% | 1.05% | 1.224% | 2.961% | 1.463% | 3.077% | 1.421% | 2.213% | 1.298% | 0.9858% | 1.582% | 3.074% | 1.762% | 2.875% | 1.35% | 1.788% | 1.706% | 1.7%
| PDM3 |  | 1.966% | 10.04% | 0.9258% | -5.02% | -1.274% | -1.534% | -0.9768% | 1.862% | 1.186% | -1.229% | 4.047% | -4.5% | 1.818% | -2.344% | -2.567% | 3.304% | -1.898% | 4.428% | 1.818% | 13.21% | -69.65%
| Z-Score (Absolute) |  | 0.29 | 2.5 | 0 | 1.63 | 0.61 | 0.67 | 0.52 | 0.26 | 0.07 | 0.59 | 0.86 | 1.49 | 0.24 | 0.9 | 0.96 | 0.65 | 0.77 | 0.96 | 0.24 | 3.37 | 19.36
| Analytical results for La in OREAS 931b (Certified Value 37.2 ppm) | [22 empty] | Certified
La | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 01; 02: 10; 03: 11; 04: 12; 05: 13; 06: 14; 07: 17; 08: 18; 09: 20; 10: 24 | ppm
|  |  | 01: PF*OES; 02: PF*OES; 03: PF*OES/MS; 04: PF*MS; 05: BF*OES; 06: PF*OES/MS; 07: PF*OES/MS; 08: PF*OES; 09: PF*MS; 10: PF*MS | 1
2
| 1 | 1 | 01: 34; 02: 20; 03: 32.8246; 04: 37.3; 05: 37; 06: 39.8; 07: 35.1; 08: 37.1; 09: 39.1; 10: 43.013 | 1
| 1 | 2 | 01: 35; 02: 28; 03: 32.2491; 04: 41; 05: 35; 06: 38.8; 07: 34.7; 08: 37.23; 09: 37.2; 10: 43.9336 | 17
| 1 | 3 | 01: 34; 02: 26; 03: 31.9757; 04: 39.3; 05: 37; 06: 38; 07: 35.7; 08: 37.33; 09: 37.9; 10: 40.8999 | 16
| 1 | 4 | 01: 36; 02: 27; 03: 32.8621; 04: 38.6; 05: 38; 06: 36.3; 07: 34.6; 08: 36.28; 09: 38.5; 10: 43.0967 | 37.1848
| 1 | 5 | 01: 36; 02: 26; 03: 34.9756; 04: 38.6; 05: 39; 06: 37; 07: 36.2; 08: 38.69; 09: 38; 10: 42.2582 | 157
| 1 | 6 | 01: 34; 02: 24; 03: 32.6336; 04: 38.9; 05: 38; 06: 36.1; 07: 35; 08: 37.7; 09: 37.7; 10: 42.9919
| Mean |  | 01: 34.8333; 02: 25.1667; 03: 32.9201; 04: 38.95; 05: 37.3333; 06: 37.6667; 07: 35.2167; 08: 37.3883; 09: 38.0667; 10: 42.6989
| Median |  | 01: 34.5; 02: 26; 03: 32.7291; 04: 38.75; 05: 37.5; 06: 37.5; 07: 35.05; 08: 37.28; 09: 37.95; 10: 43.0024
| Std Dev. |  | 01: 0.98319; 02: 2.8577; 03: 1.0641; 04: 1.2079; 05: 1.3663; 06: 1.4638; 07: 0.61779; 08: 0.79139; 09: 0.65929; 10: 1.0293
| Rel.Std.Dev. |  | 01: 2.823%; 02: 11.36%; 03: 3.232%; 04: 3.101%; 05: 3.66%; 06: 3.886%; 07: 1.754%; 08: 2.117%; 09: 1.732%; 10: 2.411%
| PDM3 |  | 01: -6.324%; 02: -32.32%; 03: -11.47%; 04: 4.747%; 05: 0.3995%; 06: 1.296%; 07: -5.293%; 08: 0.5474%; 09: 2.372%; 10: 14.83%
| Z-Score (Absolute) |  | 01: 0.91; 02: 4.4; 03: 1.6; 04: 0.57; 05: 0.01; 06: 0.11; 07: 0.77; 08: 0.01; 09: 0.25; 10: 1.93
| Analytical results for Li in OREAS 931b (Certified Value 17.8 ppm) | [22 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 10 | 12 | 14 | 15 | 16 | 17 | 18 | 20 | 22 | 24 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | [8 empty] | 1
1
| 1 | 1 | 20.3451 | 17 | 18.2 | < 10 | < 10 | 17 | < 100 | 20 | 17 | < 20 | 12 | 12 | 35.6199 | [8 empty] | 1
| 1 | 2 | 17.6489 | 17 | 18.5 | 43 | 15 | 17 | < 100 | 20 | 19 | < 20 | 11 | 14 | 18.7204 | [8 empty] | 0
| 1 | 3 | 17.9914 | 17 | 19.4 | < 10 | < 10 | 19 | < 100 | 20 | 18 | < 20 | 16 | 12 | 16.5676 | [8 empty] | 16
| 1 | 4 | 20.1272 | 17 | 18.9 | < 10 | < 10 | 18 | < 100 | 10 | 18 | < 20 | 14 | 13 | 20.2453 | [8 empty] | 17.7891
| 1 | 5 | 17.8201 | 17 | 18.4 | 58 | 20 | 18 | < 100 | 20 | 19 | < 20 | 18 | 13 | 15.8679 | [8 empty] | 158
| 1 | 6 | 18.9935 | 17 | 17.6 | 103 | 15 | 18 | < 100 | 20 | 19 | < 20 | 19 | 13 | 18.3347
| Mean |  | 01: 18.821; 02: 17; 03: 18.5; 04: 68; 05: 16.6667; 06: 17.8333; 08: 18.3333; 09: 18.3333; 12: 15; 13: 12.8333; 14: 20.8926
| Median |  | 01: 18.4924; 02: 17; 03: 18.45; 04: 58; 05: 15; 06: 18; 08: 20; 09: 18.5; 12: 15; 13: 13; 14: 18.5275
| Std Dev. |  | 01: 1.1935; 02: 0; 03: 0.61319; 04: 31.225; 05: 2.8868; 06: 0.75277; 08: 4.0825; 09: 0.8165; 12: 3.2249; 13: 0.75277; 14: 7.3828
| Rel.Std.Dev. |  | 01: 6.341%; 02: 0%; 03: 3.315%; 04: 45.92%; 05: 17.32%; 06: 4.221%; 08: 22.27%; 09: 4.454%; 12: 21.5%; 13: 5.866%; 14: 35.34%
| PDM3 |  | 01: 5.801%; 02: -4.436%; 03: 3.997%; 04: 282.3%; 05: -6.309%; 06: 0.2489%; 08: 3.06%; 09: 3.06%; 12: -15.68%; 13: -27.86%; 14: 17.45%
| Z-Score (Absolute) |  | 0.13 | 0.35 | 0.04 | 4.79 | 1.98 | 0.13 | 8.34 | 0 | 0 | 2.2 | 0.88 | 1.45 | 0.67
| Analytical results for Lu in OREAS 931b (Certified Value 0.34 ppm) | [22 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 18; 06: 20; 07: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES; 06: PF*MS; 07: PF*MS | 2
3
| 1 | 1 | 01: 0.3519; 02: 0.43; 03: 0.3; 04: 0.33; 05: 0.34; 06: 0.34; 07: < 2 | 1
| 1 | 2 | 01: 0.3628; 02: 0.52; 03: 0.29; 04: 0.31; 05: 0.35; 06: 0.33; 07: < 2 | 0
| 1 | 3 | 01: 0.3437; 02: 0.45; 03: 0.3; 04: 0.35; 05: 0.36; 06: 0.32; 07: < 2 | 16
| 1 | 4 | 01: 0.3673; 02: 0.4; 03: 0.31; 04: 0.35; 05: 0.34; 06: 0.35; 07: < 2 | 0.338
| 1 | 5 | 01: 0.3806; 02: 0.34; 03: 0.32; 04: 0.37; 05: 0.36; 06: 0.33; 07: < 2 | 159
| 1 | 6 | 01: 0.3538; 02: 0.43; 03: 0.31; 04: 0.33; 05: 0.35; 06: 0.34; 07: < 2
| Mean |  | 01: 0.36002; 02: 0.42833; 03: 0.305; 04: 0.34; 05: 0.35; 06: 0.335
| Median |  | 01: 0.3583; 02: 0.43; 03: 0.305; 04: 0.34; 05: 0.35; 06: 0.335
| Std Dev. |  | 01: 0.013078; 02: 0.059133; 03: 0.010488; 04: 0.020976; 05: 0.0089443; 06: 0.010488
| Rel.Std.Dev. |  | 01: 3.633%; 02: 13.81%; 03: 3.439%; 04: 6.169%; 05: 2.556%; 06: 3.131%
| PDM3 |  | 01: 6.513%; 02: 26.72%; 03: -9.764%; 04: 0.5907%; 05: 3.549%; 06: -0.8886%
| Z-Score (Absolute) |  | 01: 0.45; 02: 3.52; 03: 2.02; 04: 0.45; 05: 0; 06: 0.67; 07: 29.22
| Analytical results for Mg in OREAS 931b (Certified Value 0.376 wt.%) | [22 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS |  | 3
3
| 1 | 1 | 0.43 | 0.37453 | 0.37 | 0.39 | 0.38 | 0.368 | 0.368 | 0.362 | 0.38 | 0.31 | 0.38 | 0.374 | 0.34 | 0.37 | 0.38 | 0.37 | 0.39 | 0.3655 | 0.37 | 0.38993 |  | 1
| 1 | 2 | 0.44 | 0.39169 | 0.38 | 0.41 | 0.38 | 0.368 | 0.35 | 0.362 | 0.368 | 0.38 | 0.37 | 0.38 | 0.34 | 0.36 | 0.39 | 0.37 | 0.39 | 0.3649 | 0.36 | 0.40085 |  | 0
| 1 | 3 | 0.44 | 0.38367 | 0.37 | 0.42 | 0.38 | 0.374 | 0.35 | 0.362 | 0.38 | 0.35 | 0.35 | 0.38 | 0.33 | 0.38 | 0.38 | 0.37 | 0.4 | 0.3685 | 0.37 | 0.39655 |  | 16
| 1 | 4 | 0.45 | 0.37433 | 0.38 | 0.39 | 0.38 | 0.362 | 0.338 | 0.362 | 0.362 | 0.36 | 0.37 | 0.38 | 0.33 | 0.38 | 0.38 | 0.35 | 0.4 | 0.3661 | 0.38 | 0.4005 |  | 0.37557
| 1 | 5 | 0.44 | 0.38922 | 0.38 | 0.4 | 0.38 | 0.368 | 0.362 | 0.362 | 0.368 | 0.36 | 0.36 | 0.374 | 0.34 | 0.37 | 0.4 | 0.37 | 0.41 | 0.3842 | 0.37 | 0.40825 |  | 160
| 1 | 6 | 0.44 | 0.37102 | 0.38 | 0.4 | 0.38 | 0.368 | 0.368 | 0.362 | 0.368 | 0.35 | 0.38 | 0.374 | 0.33 | 0.37 | 0.39 | 0.37 | 0.4 | 0.354 | 0.38 | 0.40906
| Mean |  | 0.44 | 0.38074 | 0.37667 | 0.40167 | 0.38 | 0.368 | 0.356 | 0.362 | 0.371 | 0.35167 | 0.36833 | 0.377 | 0.335 | 0.37167 | 0.38667 | 0.36667 | 0.39833 | 0.3672 | 0.37167 | 0.40085
| Median |  | 0.44 | 0.3791 | 0.38 | 0.4 | 0.38 | 0.368 | 0.356 | 0.362 | 0.368 | 0.355 | 0.37 | 0.377 | 0.335 | 0.37 | 0.385 | 0.37 | 0.4 | 0.3658 | 0.37 | 0.40067
| Std Dev. |  | 0.0063246 | 0.0086583 | 0.005164 | 0.01169 | 6.0809e-17 | 0.0037947 | 0.012 | 6.0809e-17 | 0.0073485 | 0.023166 | 0.01169 | 0.0032863 | 0.0054772 | 0.0075277 | 0.008165 | 0.008165 | 0.0075277 | 0.0097398 | 0.0075277 | 0.0072125
| Rel.Std.Dev. |  | 1.437% | 2.274% | 1.371% | 2.91% | 1.6e-14% | 1.031% | 3.371% | 1.68e-14% | 1.981% | 6.588% | 3.174% | 0.8717% | 1.635% | 2.025% | 2.112% | 2.227% | 1.89% | 2.652% | 2.025% | 1.799%
| PDM3 |  | 17.16% | 1.377% | 0.2918% | 6.948% | 1.179% | -2.016% | -5.211% | -3.613% | -1.217% | -6.365% | -1.927% | 0.3806% | -10.8% | -1.04% | 2.954% | -2.371% | 6.061% | -2.229% | -1.04% | 6.732%
| Z-Score (Absolute) |  | 5.31 | 0.71 | 0.39 | 2.33 | 0.64 | 0.3 | 1.23 | 0.76 | 0.06 | 1.56 | 0.26 | 0.41 | 2.85 | 0 | 1.17 | 0.39 | 2.07 | 0.35 | 0 | 2.27
| Analytical results for Mn in OREAS 931b (Certified Value 0.018 wt.%) | [22 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | 3
3
| 1 | 1 | 0.0257 | 0.017003 | 0.0185 | 0.017 | 0.015 | 0.015 | 0.015 | 0.015 | 0.023 | 0.016 | 0.0174 | 0.015 | 0.0172 | 0.016 | 0.02 | 0.0172 | 0.02 | 0.0186 | 0.0165 | 0.01 | 0.01982 | 1
| 1 | 2 | 0.0256 | 0.016738 | 0.0186 | 0.018 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.0171 | 0.0174 | 0.023 | 0.0172 | 0.0158 | 0.02 | 0.017 | 0.02 | 0.017 | 0.016 | 0.02 | 0.020055 | 0
| 1 | 3 | 0.0256 | 0.017394 | 0.0189 | 0.018 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.0169 | 0.0164 | 0.023 | 0.0177 | 0.0163 | 0.02 | 0.0171 | 0.02 | 0.017 | 0.0164 | 0.02 | 0.020234 | 16
| 1 | 4 | 0.0257 | 0.01661 | 0.0192 | 0.017 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.0172 | 0.0174 | 0.023 | 0.0172 | 0.0164 | 0.02 | 0.0169 | 0.02 | 0.017 | 0.0165 | 0.02 | 0.020141 | 0.017503
| 1 | 5 | 0.0255 | 0.017712 | 0.0193 | 0.018 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.0168 | 0.0164 | 0.015 | 0.0179 | 0.0161 | 0.02 | 0.017 | 0.02 | 0.0178 | 0.0164 | 0.02 | 0.020031 | 161
| 1 | 6 | 0.0257 | 0.016432 | 0.0198 | 0.018 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.0173 | 0.017 | 0.015 | 0.0174 | 0.0163 | 0.02 | 0.0167 | 0.02 | 0.0163 | 0.0164 | 0.02 | 0.019558
| Mean |  | 0.025633 | 0.016982 | 0.01905 | 0.017667 | 0.015 | 0.015 | 0.015 | 0.015 | 0.016333 | 0.016883 | 0.017 | 0.019 | 0.017433 | 0.01615 | 0.02 | 0.016983 | 0.02 | 0.017283 | 0.016367 | 0.018333 | 0.019973
| Median |  | 0.02565 | 0.016871 | 0.01905 | 0.018 | 0.015 | 0.015 | 0.015 | 0.015 | 0.015 | 0.017 | 0.0172 | 0.019 | 0.0173 | 0.0162 | 0.02 | 0.017 | 0.02 | 0.017 | 0.0164 | 0.02 | 0.020043
| Std Dev. |  | 8.165e-05 | 0.00049088 | 0.00048477 | 0.0005164 | 0 | 0 | 0 | 0 | 0.003266 | 0.00047081 | 0.0004899 | 0.0043818 | 0.00030111 | 0.00022583 | 0 | 0.00017224 | 0 | 0.00080104 | 0.00018619 | 0.0040825 | 0.00024577
| Rel.Std.Dev. |  | 0.3185% | 2.891% | 2.545% | 2.923% | 0% | 0% | 0% | 0% | 20% | 2.789% | 2.882% | 23.06% | 1.727% | 1.398% | 0% | 1.014% | 0% | 4.635% | 1.138% | 22.27% | 1.231%
| PDM3 |  | 46.45% | -2.979% | 8.839% | 0.9354% | -14.3% | -14.3% | -14.3% | -14.3% | -6.682% | -3.54% | -2.873% | 8.553% | -0.3977% | -7.73% | 14.27% | -2.969% | 14.27% | -1.255% | -6.492% | 4.744% | 14.11%
| Z-Score (Absolute) |  | 4.37 | 0.01 | 1.04 | 0.34 | 0.76 | 0.76 | 0.76 | 0.76 | 0.11 | 0.06 | 0 | 1.19 | 0.22 | 0.43 | 1.52 | 0.01 | 1.52 | 0.15 | 0.32 | 0.67 | 1.5
| Analytical results for Mo in OREAS 931b (Certified Value 3.74 ppm) | [22 empty] | Certified
Mo | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 12: Lab | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 10; 05: 12; 06: 14; 07: 15; 08: 17; 09: 19; 10: 20; 12: 24 | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*OES; 05: PF*MS; 06: PF*OES/MS; 07: PF*MS; 08: PF*OES/MS; 09: PF*MS; 10: PF*MS; 12: PF*MS | 2
2
| 1 | 1 | 01: 3.2949; 02: 6; 03: 5; 04: < 10; 05: < 2; 06: 3; 07: 5; 08: 4; 09: 2.965; 10: 3; 12: 4.7236 | 1
| 1 | 2 | 01: 3.494; 02: 5; 03: 5; 04: < 10; 05: 3; 06: 3; 07: 5; 08: 4; 09: 3.3791; 10: 3; 12: 5.1086 | 8
| 1 | 3 | 01: 3.0097; 02: 5; 03: 6; 04: < 10; 05: 3; 06: 4; 07: 5; 08: 4; 09: 2.8383; 10: 3; 12: 3.8479 | 16
| 1 | 4 | 01: 3.3122; 02: 4; 03: 6; 04: < 10; 05: 3; 06: 3; 07: 4; 08: 4; 09: 3.2911; 10: 3; 12: 3.8861 | 3.7415
| 1 | 5 | 01: 3.5159; 02: 4; 03: 6; 04: < 10; 05: < 2; 06: 3; 07: 4; 08: 4; 09: 3.2567; 10: 3; 12: 3.9054 | 162
| 1 | 6 | 01: 2.9286; 02: 4; 03: 6; 04: < 10; 05: 3; 06: 3; 07: 5; 08: 5; 09: 3.2637; 10: 4; 12: 5.0224
| Mean |  | 01: 3.2592; 02: 4.6667; 03: 5.6667; 05: 3; 06: 3.1667; 07: 4.6667; 08: 4.1667; 09: 3.1657; 10: 3.1667; 12: 4.4156
| Median |  | 01: 3.3036; 02: 4.5; 03: 6; 05: 3; 06: 3; 07: 5; 08: 4; 09: 3.2602; 10: 3; 12: 4.3145
| Std Dev. |  | 01: 0.24358; 02: 0.8165; 03: 0.5164; 05: 0; 06: 0.40825; 07: 0.5164; 08: 0.40825; 09: 0.2129; 10: 0.40825; 12: 0.60105
| Rel.Std.Dev. |  | 01: 7.474%; 02: 17.5%; 03: 9.113%; 05: 0%; 06: 12.89%; 07: 11.07%; 08: 9.798%; 09: 6.725%; 10: 12.89%; 12: 13.61%
| PDM3 |  | 01: -12.89%; 02: 24.73%; 03: 51.45%; 05: -19.82%; 06: -15.36%; 07: 24.73%; 08: 11.36%; 09: -15.39%; 10: -15.36%; 12: 18.02%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.37; 03: 1.11; 04: 0.62; 05: 1.36; 06: 0.74; 07: 0.37; 08: 0; 09: 0.74; 10: 0.74; 12: 0.19
| Analytical results for Na in OREAS 931b (Indicative Value 0.08 wt.%) | [22 empty] | Indicative
Na | Round | Replicate | 01: Lab | 1
| No | No | 01: 13 | wt.%
|  |  | 01: BF*OES | 3
3
| 1 | 1 | 01: 0.089 | 1
| 1 | 2 | 01: 0.082 | 19
| 1 | 3 | 01: 0.074 | 16
| 1 | 4 | 01: 0.082 | 0.080366
| 1 | 5 | 01: 0.082 | 25
| 1 | 6 | 01: 0.074
| Mean |  | 01: 0.0805
| Median |  | 01: 0.082
| Std Dev. |  | 01: 0.0057184
| Rel.Std.Dev. |  | 01: 7.104%
| PDM3 |  | 01: 0.1668%
Z-Score (Absolute) | bar: 01
| Analytical results for Nb in OREAS 931b (Certified Value 8.37 ppm) | [22 empty] | Certified
Nb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 12: Lab; 13: Lab | 1
| No | No | 01: 01; 02: 02; 03: 10; 04: 12; 05: 13; 06: 14; 07: 17; 08: 18; 09: 19; 10: 20; 12: 22; 13: 24 | ppm
|  |  | 01: PF*OES; 02: PF*MS; 03: PF*OES; 04: PF*MS; 05: BF*OES/MS; 06: PF*OES/MS; 07: PF*OES/MS; 08: PF*OES; 09: PF*MS; 10: PF*MS; 12: PF*OES/MS; 13: PF*MS | 2
2
| 1 | 1 | 01: 19; 02: 7.9629; 03: 15; 04: 8; 05: 8.7; 06: 10; 07: 9.6; 08: < 20; 09: 9.9876; 10: 8; 12: 21; 13: 5.935 | 1
| 1 | 2 | 01: 20; 02: 8.657; 03: 17; 04: 9; 05: 8.3; 06: 10; 07: 9; 08: < 20; 09: 11.2143; 10: 8; 12: 21; 13: 5.9099 | 10
| 1 | 3 | 01: 18; 02: 8.1238; 03: 20; 04: 9; 05: 8.9; 06: 10; 07: 7.7; 08: < 20; 09: 9.3056; 10: 8; 12: 20; 13: 5.8394 | 16
| 1 | 4 | 01: 19; 02: 7.6395; 03: 22; 04: 8; 05: 8.4; 06: 10; 07: 8.8; 08: < 20; 09: 10.4937; 10: 8; 12: 19; 13: 6.4284 | 8.3708
| 1 | 5 | 01: 19; 02: 7.6484; 03: 18; 04: 6; 05: 8.2; 06: 10; 07: 9.5; 08: < 20; 09: 10.905; 10: 8; 12: 21; 13: 16.3012 | 163
| 1 | 6 | 01: 18; 02: 7.1021; 03: 23; 04: 6; 05: 8.5; 06: 10; 07: 8.4; 08: < 20; 09: 9.5917; 10: 7; 12: 19; 13: 16.4825
| Mean |  | 01: 18.8333; 02: 7.8556; 03: 19.1667; 04: 7.6667; 05: 8.5; 06: 10; 07: 8.8333; 09: 10.2496; 10: 7.8333; 12: 20.1667; 13: 9.4827
| Median |  | 01: 19; 02: 7.8057; 03: 19; 04: 8; 05: 8.45; 06: 10; 07: 8.9; 09: 10.2406; 10: 8; 12: 20.5; 13: 6.1817
| Std Dev. |  | 01: 0.75277; 02: 0.52594; 03: 3.0605; 04: 1.3663; 05: 0.26077; 06: 0; 07: 0.71181; 09: 0.75; 10: 0.40825; 12: 0.98319; 13: 5.3562
| Rel.Std.Dev. |  | 01: 3.997%; 02: 6.695%; 03: 15.97%; 04: 17.82%; 05: 3.068%; 06: 0%; 07: 8.058%; 09: 7.317%; 10: 5.212%; 12: 4.875%; 13: 56.48%
| PDM3 |  | 01: 125%; 02: -6.155%; 03: 129%; 04: -8.412%; 05: 1.543%; 06: 19.46%; 07: 5.525%; 09: 22.44%; 10: -6.421%; 12: 140.9%; 13: 13.28%
| Z-Score (Absolute) |  | 01: 3.92; 02: 0.81; 03: 4.06; 04: 0.89; 05: 0.54; 06: 0.11; 07: 0.39; 08: 0.11; 09: 0.22; 10: 0.82; 12: 4.5; 13: 0.11
| Analytical results for Nd in OREAS 931b (Certified Value 31.6 ppm) | [22 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS | 1
2
| 1 | 1 | 01: 31.2702; 02: 29.5; 03: 28.6; 04: 29.6; 05: 32.2; 06: 32.1; 07: 33.9; 08: 36.0704 | 1
| 1 | 2 | 01: 31.1043; 02: 31.4; 03: 28.5; 04: 29.1; 05: 31.6; 06: 32.1; 07: 32.1; 08: 39.243 | 5
| 1 | 3 | 01: 30.5511; 02: 31.8; 03: 27.5; 04: 32.5; 05: 32.3; 06: 32.3; 07: 32.5; 08: 37.3463 | 16
| 1 | 4 | 01: 31.7539; 02: 29.8; 03: 28; 04: 32.6; 05: 29; 06: 31; 07: 33; 08: 35.9929 | 31.6456
| 1 | 5 | 01: 33.0116; 02: 30.7; 03: 28; 04: 31.9; 05: 32.2; 06: 32.5; 07: 32.8; 08: 39.1137 | 164
| 1 | 6 | 01: 30.8881; 02: 31; 03: 27.9; 04: 31.5; 05: 31.5; 06: 32.7; 07: 32.4; 08: 38.0791
| Mean |  | 01: 31.4299; 02: 30.7; 03: 28.0833; 04: 31.2; 05: 31.4667; 06: 32.1167; 07: 32.7833; 08: 37.6409
| Median |  | 01: 31.1872; 02: 30.85; 03: 28; 04: 31.7; 05: 31.9; 06: 32.2; 07: 32.65; 08: 37.7127
| Std Dev. |  | 01: 0.87225; 02: 0.89889; 03: 0.40702; 04: 1.4967; 05: 1.2549; 06: 0.5947; 07: 0.63061; 08: 1.4281
| Rel.Std.Dev. |  | 01: 2.775%; 02: 2.928%; 03: 1.449%; 04: 4.797%; 05: 3.988%; 06: 1.852%; 07: 1.924%; 08: 3.794%
| PDM3 |  | 01: -0.6817%; 02: -2.988%; 03: -11.26%; 04: -1.408%; 05: -0.5654%; 06: 1.489%; 07: 3.595%; 08: 18.95%
| Z-Score (Absolute) |  | 01: 0.02; 02: 0.71; 03: 3.2; 04: 0.24; 05: 0.02; 06: 0.64; 07: 1.27; 08: 5.9
| Analytical results for Ni in OREAS 931b (Indicative Value 21.9 ppm) | [22 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 |  | ppm
|  |  | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES | PF*MS |  | 1
1
| 1 | 1 | < 10 | < 20 | < 20 | < 20 | < 20 | 70 | < 20 | 20 | 21 | 29 | 77 | 12 | 30 | < 50 | 30 | < 50 | 21.0898 | 26 | < 100 | 20.4292 |  | 1
| 1 | 2 | < 10 | < 20 | < 20 | 20 | < 20 | 70 | < 20 | 30 | 14 | 23 | < 20 | 10 | 26 | < 50 | 30 | 80 | 25.9565 | 23 | < 100 | 19.7987 |  | 12
| 1 | 3 | < 10 | < 20 | < 20 | 40 | < 20 | 60 | 20 | < 20 | 21 | 23 | 27 | 11 | 27 | < 50 | 20 | 60 | 18.8523 | 29 | < 100 | 56.0037 |  | 16
| 1 | 4 | < 10 | < 20 | < 20 | 20 | < 20 | 70 | < 20 | 30 | 14 | 25 | 24 | 12 | 24 | < 50 | 20 | < 50 | 34.6673 | 26 | < 100 | 10.5928 |  | 21.8668
| 1 | 5 | < 10 | < 20 | < 20 | 20 | < 20 | 80 | 50 | < 20 | 16 | 21 | < 20 | 13 | 22 | < 50 | 20 | < 50 | 25.4452 | 31 | < 100 | 19.4204 |  | 26
| 1 | 6 | < 10 | < 20 | < 20 | < 20 | 20 | 100 | 20 | 20 | 14 | 17 | < 20 | 10 | 28 | < 50 | 40 | < 50 | 20.9227 | 24 | < 100 | 17.655
| Mean |  |  |  |  | 25 | 20 | 75 | 30 | 25 | 16.6667 | 23 | 42.6667 | 11.3333 | 26.1667 |  | 26.6667 | 70 | 24.489 | 26.5 |  | 23.9833
| Median |  |  |  |  | 20 | 20 | 70 | 20 | 25 | 15 | 23 | 27 | 11.5 | 26.5 |  | 25 | 70 | 23.2675 | 26 |  | 19.6095
| Std Dev. |  |  |  |  | 10 |  | 13.784 | 17.3205 | 5.7735 | 3.4448 | 4 | 29.7714 | 1.2111 | 2.8577 |  | 8.165 | 14.1421 | 5.7046 | 3.0166 |  | 16.0973
| Rel.Std.Dev. |  |  |  |  | 40% |  | 18.38% | 57.74% | 23.09% | 20.67% | 17.39% | 69.78% | 10.69% | 10.92% |  | 30.62% | 20.2% | 23.29% | 11.38% |  | 67.12%
| PDM3 |  |  |  |  | 14.33% | -8.537% | 243% | 37.19% | 14.33% | -23.78% | 5.182% | 95.12% | -48.17% | 19.66% |  | 21.95% | 220.1% | 11.99% | 21.19% |  | 9.679%
| Z-Score (Absolute) |  | 3.57 | 2.61 | 2.61 | 0.67 | 2.28 | 9.95 | 0.67 | 0.67 | 1.32 | 0.09 | 0.55 | 2.35 | 0.52 | 0.29 | 0.61 | 3.19 | 0.19 | 0.58 | 5.12 | 0.09
| Analytical results for P in OREAS 931b (Certified Value 0.03 wt.%) | [22 empty] | Certified
P | Round | Replicate | 01: Lab; 02: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 12: Lab |  | 1
| No | No | 01: 01; 02: 04; 04: 10; 05: 12; 06: 13; 07: 14; 08: 15; 09: 19; 10: 20; 12: 24 |  | wt.%
|  |  | 01: PF*OES; 02: PF*OES/MS; 04: PF*OES; 05: PF*OES; 06: BF*OES/MS; 07: PF*OES/MS; 08: PF*OES; 09: PF*OES; 10: PF*OES; 12: PF*MS |  | 3
3
| 1 | 1 | 01: 0.1309; 02: 0.03; 04: < 0.01; 05: 0.04; 06: 0.026; 07: 0.2; 08: 0.03; 09: < 0.05; 10: 0.03; 12: 0.028272 |  | 1
| 1 | 2 | 01: 0.1348; 02: 0.03; 04: < 0.01; 05: 0.04; 06: 0.031; 07: 0.2; 08: 0.03; 09: < 0.05; 10: 0.03; 12: 0.027543 |  | 0
| 1 | 3 | 01: 0.1314; 02: 0.03; 04: < 0.01; 05: 0.02; 06: 0.026; 07: 0.21; 08: 0.03; 09: < 0.05; 10: 0.03; 12: 0.027719 |  | 16
| 1 | 4 | 01: 0.1322; 02: 0.03; 04: < 0.01; 05: 0.03; 06: 0.026; 07: 0.21; 08: 0.03; 09: < 0.05; 10: 0.03; 12: 0.031164 |  | 0.029566
| 1 | 5 | 01: 0.1344; 02: 0.03; 04: < 0.01; 05: 0.04; 06: 0.031; 07: 0.22; 08: 0.03; 09: < 0.05; 10: 0.03; 12: 0.028162 |  | 165
| 1 | 6 | 01: 0.1315; 02: 0.03; 04: < 0.01; 05: 0.03; 06: 0.026; 07: 0.22; 08: 0.02; 09: < 0.05; 10: 0.03; 12: 0.02876
| Mean |  | 01: 0.13253; 02: 0.03; 05: 0.033333; 06: 0.027667; 07: 0.21; 08: 0.028333; 10: 0.03; 12: 0.028603
| Median |  | 01: 0.13185; 02: 0.03; 05: 0.035; 06: 0.026; 07: 0.21; 08: 0.03; 10: 0.03; 12: 0.028217
| Std Dev. |  | 01: 0.0016585; 02: 0; 05: 0.008165; 06: 0.002582; 07: 0.0089443; 08: 0.0040825; 10: 0; 12: 0.0013264
| Rel.Std.Dev. |  | 01: 1.251%; 02: 0%; 05: 24.49%; 06: 9.332%; 07: 4.259%; 08: 14.41%; 10: 0%; 12: 4.637%
| PDM3 |  | 01: 348.3%; 02: 1.469%; 05: 12.74%; 06: -6.423%; 07: 610.3%; 08: -4.169%; 10: 1.469%; 12: -3.256%
| Z-Score (Absolute) |  | 01: 24.44; 02: 0.17; 04: 5.75; 05: 0.95; 06: 0.39; 07: 42.78; 08: 0.23; 09: 1.02; 10: 0.17; 12: 0.17
| Analytical results for Pb in OREAS 931b (Certified Value 121 ppm) | [22 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 |  | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES |  | 0
0
| 1 | 1 | 73 | 100.8079 | 129 | 146 | 100 | 100 | 100 | 100 | 100 | 93 | 113 | 129 | 141 | < 300 | 130 | 100 | 136.839 | 125 | 84 | 100 |  | 1
| 1 | 2 | 74 | 103.1732 | 126 | 149 | 100 | 100 | 100 | 100 | 100 | 125 | 111 | 127 | 147 | < 300 | 145 | 100 | 124.5907 | 120 | 64 | < 100 |  | 13
| 1 | 3 | 74 | 95.3332 | 126 | 145 | 100 | 100 | 100 | 100 | 100 | 107 | 110 | 124 | 145 | < 300 | 122 | 100 | 159.5862 | 122 | 93 | 200 |  | 16
| 1 | 4 | 75 | 95.1766 | 126 | 139 | 100 | 100 | 200 | 100 | 100 | 114 | 119 | 123 | 138 | < 300 | 128 | 100 | 147.2756 | 124 | 85 | 100 |  | 120.954
| 1 | 5 | 76 | 98.6227 | 128 | 151 | 100 | 100 | 100 | 100 | 100 | 115 | 103 | 126 | 144 | < 300 | 128 | 100 | 123.1137 | 123 | 91 | 200 |  | 166
| 1 | 6 | 74 | 92.2294 | 129 | 148 | 100 | 100 | 100 | < 100 | 200 | 101 | 118 | 122 | 137 | < 300 | 131 | 100 | 145.4151 | 123 | 66 | 200
| Mean |  | 74.3333 | 97.5571 | 127.3333 | 146.3333 | 100 | 100 | 116.6667 | 100 | 116.6667 | 109.1667 | 112.3333 | 125.1667 | 142 |  | 130.6667 | 100 | 139.4701 | 122.8333 | 80.5 | 160
| Median |  | 74 | 96.9779 | 127 | 147 | 100 | 100 | 100 | 100 | 100 | 110.5 | 112 | 125 | 142.5 |  | 129 | 100 | 141.1271 | 123 | 84.5 | 200
| Std Dev. |  | 1.0328 | 4.0556 | 1.5055 | 4.1793 | 0 | 0 | 40.8248 | 0 | 40.8248 | 11.3211 | 5.8538 | 2.6394 | 4 |  | 7.6855 | 0 | 14.1193 | 1.7224 | 12.502 | 54.7723
| Rel.Std.Dev. |  | 1.389% | 4.157% | 1.182% | 2.856% | 0% | 0% | 34.99% | 0% | 34.99% | 10.37% | 5.211% | 2.109% | 2.817% |  | 5.882% | 0% | 10.12% | 1.402% | 15.53% | 34.23%
| PDM3 |  | -38.54% | -19.34% | 5.274% | 20.98% | -17.32% | -17.32% | -3.545% | -17.32% | -3.545% | -9.745% | -7.127% | 3.483% | 17.4% |  | 8.03% | -17.32% | 15.31% | 1.554% | -33.45% | 32.28%
| Z-Score (Absolute) |  | 2.14 | 1.16 | 0.09 | 0.89 | NA | NA | NA | NA | NA | 0.67 | 0.54 | 0 | 0.71 | 1.05 | 0.23 | NA | 0.6 | 0.1 | 1.88 | NA
Results from laboratories 5, 6, 7, 8, 9, 18 and 23 were removed due to their 100 ppm reading resolution.
| Analytical results for Pr in OREAS 931b (Certified Value 8.45 ppm) | [22 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS |  | 2
3
| 1 | 1 | 01: 9.0287; 02: 8.47; 03: 8.08; 04: 7.6; 05: 8.3; 06: 8.49; 07: 8.9; 08: 9.7319 |  | 1
| 1 | 2 | 01: 8.7979; 02: 8.91; 03: 8.02; 04: 7.54; 05: 8.4; 06: 8.36; 07: 8.5; 08: 10.7659 |  | 14
| 1 | 3 | 01: 8.8771; 02: 8.46; 03: 7.91; 04: 8.55; 05: 8.7; 06: 8.46; 07: 8.69; 08: 10.1076 |  | 16
| 1 | 4 | 01: 8.8154; 02: 8.51; 03: 8.24; 04: 8.5; 05: 8; 06: 8.35; 07: 8.79; 08: 9.815 |  | 8.4529
| 1 | 5 | 01: 9.5558; 02: 8.17; 03: 7.85; 04: 8.33; 05: 8.8; 06: 8.79; 07: 8.67; 08: 10.5897 |  | 167
| 1 | 6 | 01: 9.072; 02: 8.45; 03: 8.03; 04: 8.45; 05: 8.6; 06: 8.64; 07: 8.5; 08: 10.1408
| Mean |  | 01: 9.0245; 02: 8.495; 03: 8.0217; 04: 8.1617; 05: 8.4667; 06: 8.515; 07: 8.675; 08: 10.1918
| Median |  | 01: 8.9529; 02: 8.465; 03: 8.025; 04: 8.39; 05: 8.5; 06: 8.475; 07: 8.68; 08: 10.1242
| Std Dev. |  | 01: 0.28332; 02: 0.23747; 03: 0.13644; 04: 0.46447; 05: 0.29439; 06: 0.17097; 07: 0.15834; 08: 0.41258
| Rel.Std.Dev. |  | 01: 3.139%; 02: 2.795%; 03: 1.701%; 04: 5.691%; 05: 3.477%; 06: 2.008%; 07: 1.825%; 08: 4.048%
| PDM3 |  | 01: 6.762%; 02: 0.4982%; 03: -5.101%; 04: -3.445%; 05: 0.163%; 06: 0.7348%; 07: 2.628%; 08: 20.57%
| Z-Score (Absolute) |  | 01: 1.36; 02: 0.03; 03: 1.27; 04: 0.9; 05: 0.1; 06: 0.03; 07: 0.45; 08: 4.43
| Analytical results for Rb in OREAS 931b (Certified Value 124 ppm) | [22 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [9 empty] | 0
0
| 1 | 1 | 134.2536 | 125.9 | 113 | 137 | 119.9 | 122 | 126.2 | 125 | 123 | 119.2877 | 123 | 181.4944 | [9 empty] | 1
| 1 | 2 | 133.4707 | 123.5 | 124 | 139 | 109.4 | 117 | 124.3 | 123 | 122 | 118.0911 | 120 | 192.0447 | [9 empty] | 15
| 1 | 3 | 127.5209 | 126.2 | 119 | 134 | 118.8 | 115 | 127.8 | 130 | 124 | 114.9996 | 120 | 173.7705 | [9 empty] | 16
| 1 | 4 | 134.4327 | 121.8 | 124 | 130 | 119.5 | 113 | 124.1 | 132 | 119 | 117.4475 | 120 | 161.0367 | [9 empty] | 123.8695
| 1 | 5 | 133.1896 | 124 | 124 | 134 | 115.7 | 121 | 125.8 | 124 | 123 | 114.9857 | 120 | 201.8042 | [9 empty] | 168
| 1 | 6 | 128.0906 | 121.5 | 127 | 134 | 117.9 | 121 | 128.7 | 124 | 125 | 115.0564 | 119 | 173.6575
| Mean |  | 131.8263 | 123.8167 | 121.8333 | 134.6667 | 116.8667 | 118.1667 | 126.15 | 126.3333 | 122.6667 | 116.6447 | 120.3333 | 180.6347
| Median |  | 133.3301 | 123.75 | 124 | 134 | 118.35 | 119 | 126 | 124.5 | 123 | 116.252 | 120 | 177.6325
| Std Dev. |  | 3.154 | 1.9793 | 5.0365 | 3.0768 | 3.951 | 3.7103 | 1.8426 | 3.7238 | 2.0656 | 1.8816 | 1.3663 | 14.553
| Rel.Std.Dev. |  | 2.393% | 1.599% | 4.134% | 2.285% | 3.381% | 3.14% | 1.461% | 2.948% | 1.684% | 1.613% | 1.135% | 8.057%
| PDM3 |  | 6.424% | -0.04264% | -1.644% | 8.717% | -5.653% | -4.604% | 1.841% | 1.989% | -0.971% | -5.833% | -2.855% | 45.83%
| Z-Score (Absolute) |  | 1.42 | 0.09 | 0.23 | 1.89 | 1.05 | 0.84 | 0.48 | 0.51 | 0.09 | 1.09 | 0.48 | 9.48
| Analytical results for Re in OREAS 931b (Certified Value < 0.1 ppm) | [22 empty] | Certified
Re | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 12: Lab; 13: Lab; 14: Lab | [9 empty] | 1
| No | No | 02: 02; 03: 03; 04: 04; 12: 12; 13: 15; 14: 24 | [9 empty] | ppm
|  |  | 02: PF*MS; 03: PF*MS; 04: PF*OES/MS; 12: PF*MS; 13: PF*MS; 14: PF*MS | [9 empty] | 2
2
| 1 | 1 | 02: < 0.1; 03: < 0.1; 04: < 0.05; 12: < 0.05; 13: < 0.1; 14: < 0.02 | [9 empty] | 1
| 1 | 2 | 02: < 0.1; 03: < 0.1; 04: < 0.05; 12: < 0.05; 13: < 0.1; 14: < 0.02 | [9 empty] | 0
| 1 | 3 | 02: < 0.1; 03: < 0.1; 04: < 0.05; 12: < 0.05; 13: < 0.1; 14: < 0.02 | [9 empty] | 16
| 1 | 4 | 02: < 0.1; 03: < 0.1; 04: < 0.05; 12: < 0.05; 13: < 0.1; 14: < 0.02 | [9 empty] | < 0.1
| 1 | 5 | 02: < 0.1; 03: < 0.1; 04: < 0.05; 12: < 0.05; 13: < 0.1; 14: < 0.02 | [9 empty] | 169
| 1 | 6 | 02: < 0.1; 03: < 0.1; 04: < 0.05; 12: < 0.05; 13: < 0.1; 14: < 0.02
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 02: 0.67; 03: 0.67; 04: 0.67; 12: 0.67; 13: 0.67; 14: 1.48
| Analytical results for S in OREAS 931b (Certified Value 4.42 wt.%) | [22 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 12 | 14 | 15 | 16 | 17 | 18 | 24 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | [5 empty] | 2
3
| 1 | 1 | 4.43 | 4.7201 | 4.06 | 4.29 | 4.52 | 4.53 | 4.68 | 4.58 | 4.41 | 4.13 | 4.65 | 4.22 | 4.36 | 4.34 | 4.35 | 3.4808 | [5 empty] | 1
| 1 | 2 | 4.42 | 4.5411 | 4.33 | 4.49 | 4.51 | 4.49 | 4.57 | 4.6 | 4.27 | 4.12 | 4.74 | 4.17 | 4.41 | 4.38 | 4.44 | 3.5498 | [5 empty] | 0
| 1 | 3 | 4.44 | 4.4688 | 4 | 4.32 | 4.54 | 4.51 | 4.5 | 4.6 | 4.37 | 4.02 | 4.53 | 4.19 | 4.38 | 4.29 | 4.52 | 3.0979 | [5 empty] | 16
| 1 | 4 | 4.47 | 4.5272 | 4.31 | 4.39 | 4.51 | 4.5 | 4.35 | 4.59 | 4.32 | 4.12 | 4.55 | 4.2 | 4.35 | 4.46 | 4.54 | 3.7199 | [5 empty] | 4.4236
| 1 | 5 | 4.44 | 4.5906 | 4.41 | 4.53 | 4.56 | 4.44 | 4.67 | 4.64 | 4.36 | 4 | 4.64 | 4.16 | 4.46 | 4.38 | 4.54 | 4.3103 | [5 empty] | 170
| 1 | 6 | 4.47 | 4.7604 | 4.38 | 4.4 | 4.53 | 4.49 | 4.62 | 4.65 | 4.32 | 4.12 | 4.57 | 4.19 | 4.44 | 4.26 | 4.49 | 4.3172
| Mean |  | 4.445 | 4.6014 | 4.2483 | 4.4033 | 4.5283 | 4.4933 | 4.565 | 4.61 | 4.3417 | 4.085 | 4.6133 | 4.1883 | 4.4 | 4.3517 | 4.48 | 3.746
| Median |  | 4.44 | 4.5658 | 4.32 | 4.395 | 4.525 | 4.495 | 4.595 | 4.6 | 4.34 | 4.12 | 4.605 | 4.19 | 4.395 | 4.36 | 4.505 | 3.6349
| Std Dev. |  | 0.020736 | 0.11506 | 0.17383 | 0.09331 | 0.019408 | 0.030111 | 0.1247 | 0.028284 | 0.048751 | 0.058566 | 0.078655 | 0.02137 | 0.044272 | 0.071671 | 0.074027 | 0.48461
| Rel.Std.Dev. |  | 0.4665% | 2.501% | 4.092% | 2.119% | 0.4286% | 0.6701% | 2.732% | 0.6135% | 1.123% | 1.434% | 1.705% | 0.5102% | 1.006% | 1.647% | 1.652% | 12.94%
| PDM3 |  | 0.4827% | 4.017% | -3.963% | -0.4592% | 2.367% | 1.575% | 3.195% | 4.213% | -1.853% | -7.655% | 4.288% | -5.319% | -0.5345% | -1.627% | 1.274% | -15.32%
| Z-Score (Absolute) |  | 0.11 | 0.98 | 0.97 | 0.11 | 0.57 | 0.38 | 0.78 | 1.02 | 0.45 | 1.87 | 1.04 | 1.3 | 0.13 | 0.4 | 0.31 | 3.73
| Analytical results for Sb in OREAS 931b (Certified Value 3.99 ppm) | [22 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 10 | 12 | 14 | 15 | 17 | 19 | 20 | 23 | [10 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*MS | PF*OES | [10 empty] | 2
2
| 1 | 1 | < 10 | 2.8483 | 4.1 | < 50 | 3 | 4.7 | 4.8 | 4 | 3.4544 | 4.1 | < 100 | [10 empty] | 1
| 1 | 2 | < 10 | 2.8022 | 4.3 | < 50 | 4 | 4.8 | 4.7 | 4 | 3.4556 | 4.1 | < 100 | [10 empty] | 16
| 1 | 3 | < 10 | 2.5376 | 4.2 | < 50 | 4 | 5 | 4.4 | 3 | 3.5338 | 4 | < 100 | [10 empty] | 16
| 1 | 4 | < 10 | 2.7023 | 4.4 | < 50 | 3 | 4.4 | 4.4 | 4 | 3.6322 | 4.2 | < 100 | [10 empty] | 3.9898
| 1 | 5 | < 10 | 3.1275 | 4.3 | < 50 | 3 | 4.8 | 4.7 | 4 | 3.442 | 4.1 | < 100 | [10 empty] | 171
| 1 | 6 | < 10 | 2.3078 | 4.1 | < 50 | 455 | 4.8 | 4.6 | 4 | 3.4704 | 3.9 | < 100
| Mean |  | 02: 2.721; 03: 4.2333; 05: 78.6667; 06: 4.75; 07: 4.6; 08: 3.8333; 09: 3.4981; 12: 4.0667
| Median |  | 02: 2.7523; 03: 4.25; 05: 3.5; 06: 4.8; 07: 4.65; 08: 4; 09: 3.463; 12: 4.1
| Std Dev. |  | 02: 0.28027; 03: 0.12111; 05: 184.3656; 06: 0.19748; 07: 0.16733; 08: 0.40825; 09: 0.07334; 12: 0.10328
| Rel.Std.Dev. |  | 02: 10.3%; 03: 2.861%; 05: 234.4%; 06: 4.158%; 07: 3.638%; 08: 10.65%; 09: 2.097%; 12: 2.54%
| PDM3 |  | 02: -31.8%; 03: 6.103%; 05: 1872%; 06: 19.05%; 07: 15.29%; 08: -3.923%; 09: -12.33%; 12: 1.926%
| Z-Score (Absolute) |  | 0.51 | 2.38 | 0.46 | 25.79 | NA | 0.19 | 0 | NA | 1.39 | 0.67 | 57.4
Results from laboratories 12 and 17 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 931b (Certified Value 9.44 ppm) | [22 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 10 | 13 | 14 | 15 | 19 | 20 | 24 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | [10 empty] | 2
2
| 1 | 1 | 9 | < 20 | < 10 | 7 | < 5 | 10 | 10 | < 10 | 54.0071 | 10 | 10.5567 | [10 empty] | 1
| 1 | 2 | 10 | < 20 | < 10 | 7 | 7 | 9 | 10 | < 10 | 59.8603 | 9 | 10.7434 | [10 empty] | 0
| 1 | 3 | 9 | < 20 | < 10 | 7 | < 5 | 10 | 10 | < 10 | 28.398 | 10 | 10.4718 | [10 empty] | 16
| 1 | 4 | 10 | < 20 | < 10 | 7 | 6 | 10 | 9 | < 10 | 37.7348 | 10 | 10.7334 | [10 empty] | 9.4436
| 1 | 5 | 9 | < 20 | < 10 | 7 | 6 | 10 | 10 | < 10 | 56.3071 | 10 | 10.6892 | [10 empty] | 172
| 1 | 6 | 9 | < 20 | < 10 | 7 | < 5 | 10 | 10 | < 10 | 39.1693 | 10 | 10.7744
| Mean |  | 01: 9.3333; 04: 7; 05: 6.3333; 06: 9.8333; 07: 9.8333; 09: 45.9128; 12: 9.8333; 14: 10.6615
| Median |  | 01: 9; 04: 7; 05: 6; 06: 10; 07: 10; 09: 46.5882; 12: 10; 14: 10.7113
| Std Dev. |  | 01: 0.5164; 04: 0; 05: 0.57735; 06: 0.40825; 07: 0.40825; 09: 12.5476; 12: 0.40825; 14: 0.12031
| Rel.Std.Dev. |  | 01: 5.533%; 04: 0%; 05: 9.116%; 06: 4.152%; 07: 4.152%; 09: 27.33%; 12: 4.152%; 14: 1.128%
| PDM3 |  | 01: -1.167%; 04: -25.88%; 05: -32.94%; 06: 4.127%; 07: 4.127%; 09: 386.2%; 12: 4.127%; 14: 12.9%
| Z-Score (Absolute) |  | 0.41 | 0.14 | 3.94 | 2.31 | 4.41 | 0 | 0 | 3.94 | 29.38 | 0 | 0.67
| Analytical results for Se in OREAS 931b (Certified Value 46.2 ppm) | [22 empty] | Certified
Se | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [5 empty] | 1
| No | No | 01: 02; 02: 03; 03: 14; 04: 15; 05: 17; 06: 19 | [5 empty] | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS | [5 empty] | 1
2
| 1 | 1 | 01: 44.9464; 02: 44; 03: 42; 04: 48; 05: 48; 06: 4.059 | [5 empty] | 1
| 1 | 2 | 01: 45.1253; 02: 43; 03: 38.8; 04: 45; 05: 42; 06: 2.911 | [5 empty] | 0
| 1 | 3 | 01: 46.833; 02: 44; 03: 58.6; 04: 46; 05: 48; 06: 3.3104 | [5 empty] | 16
| 1 | 4 | 01: 46.2622; 02: 45; 03: 52; 04: 45; 05: 44; 06: 2.8061 | [5 empty] | 46.2437
| 1 | 5 | 01: 44.7573; 02: 44; 03: 53.9; 04: 47; 05: 44; 06: 3.4583 | [5 empty] | 173
| 1 | 6 | 01: 47.6871; 02: 48; 03: 55.4; 04: 48; 05: 42; 06: 2.556
| Mean |  | 01: 45.9352; 02: 44.6667; 03: 50.1167; 04: 46.5; 05: 44.6667; 06: 3.1835
| Median |  | 01: 45.6938; 02: 44; 03: 52.95; 04: 46.5; 05: 44; 06: 3.1107
| Std Dev. |  | 01: 1.1835; 02: 1.7512; 03: 7.8954; 04: 1.3784; 05: 2.7325; 06: 0.54195
| Rel.Std.Dev. |  | 01: 2.576%; 02: 3.921%; 03: 15.75%; 04: 2.964%; 05: 6.118%; 06: 17.02%
| PDM3 |  | 01: -0.6671%; 02: -3.41%; 03: 8.375%; 04: 0.5542%; 05: -3.41%; 06: -93.12%
| Z-Score (Absolute) |  | 01: 0.47; 02: 0.47; 03: 3.54; 04: 0.88; 05: 0.47; 06: 30.98
| Analytical results for Si in OREAS 931b (Certified Value 29.73 wt.%) | [22 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS |  |  |  |  | 2
3
| 1 | 1 | 30.64 | 30.9947 | 29.3 | 28.98 | 29.31 | 30.99 | 29.31 | 29.59 | 28.2 | 30.694 | 30.3 | 29.6 | 29.4 | 28.5 | 29.1908 | 29.6 | 14.5326 |  |  |  |  | 1
| 1 | 2 | 31.23 | 31.4689 | 30.1 | 28.98 | 29.5 | 30.2 | 29.4 | 29.4 | 28.7 | 30.39 | 31.1 | 29.5 | > 30 | 28.8 | 29.3166 | 28.7 | 14.6475 |  |  |  |  | 0
| 1 | 3 | 30.63 | 31.0198 | 29.7 | 29.12 | 29.68 | 30.01 | 29.5 | 29.31 | 28.2 | 30.357 | 31.3 | 30 | 25.2 | 29.1 | 29.5265 | 29.5 | 14.2682 |  |  |  |  | 16
| 1 | 4 | 31.11 | 31.5111 | 29.8 | 28.7 | 29.31 | 28.89 | 29.4 | 29.5 | 28.3 | 30.366 | 32.4 | 30.7 | 28.2 | 29.2 | 29.3025 | > 30 | 15.2916 |  |  |  |  | 29.7281
| 1 | 5 | 31.34 | 30.9799 | 29.6 | 29.22 | 29.22 | 30.39 | 30.1 | 29.68 | 27.8 | 30.352 | 32 | 29.8 | 28.9 | 29.5 | 29.8205 | 29.9 | 23.8128 |  |  |  |  | 174
| 1 | 6 | 31.42 | 31.4245 | 30.2 | 28.98 | 29.5 | 30.99 | 29.78 | 29.5 | 28.1 | 30.329 | 32.1 | 29.9 | 26.4 | 29.2 | 29.2142 | 29.2 | 23.8934
| Mean |  | 31.0617 | 31.2331 | 29.7833 | 28.9967 | 29.42 | 30.245 | 29.5817 | 29.4967 | 28.2167 | 30.4147 | 31.5333 | 29.9167 | 27.62 | 29.05 | 29.3952 | 29.38 | 17.741
| Median |  | 31.17 | 31.2221 | 29.75 | 28.98 | 29.405 | 30.295 | 29.45 | 29.5 | 28.2 | 30.3615 | 31.65 | 29.85 | 28.2 | 29.15 | 29.3096 | 29.5 | 14.9695
| Std Dev. |  | 0.34661 | 0.25923 | 0.33116 | 0.17546 | 0.17006 | 0.77773 | 0.30149 | 0.13155 | 0.29269 | 0.13827 | 0.78145 | 0.42622 | 1.7669 | 0.35071 | 0.23976 | 0.45497 | 4.7464
| Rel.Std.Dev. |  | 1.116% | 0.83% | 1.112% | 0.6051% | 0.578% | 2.571% | 1.019% | 0.446% | 1.037% | 0.4546% | 2.478% | 1.425% | 6.397% | 1.207% | 0.8157% | 1.549% | 26.75%
| PDM3 |  | 4.486% | 5.063% | 0.1857% | -2.46% | -1.036% | 1.739% | -0.4926% | -0.7785% | -5.084% | 2.309% | 6.072% | 0.6343% | -7.091% | -2.281% | -1.12% | -1.171% | -40.32%
| Z-Score (Absolute) |  | 2.12 | 2.35 | 0.39 | 0.67 | 0.11 | 1.01 | 0.12 | 0 | 1.73 | 1.24 | 2.75 | 0.57 | 2.54 | 0.6 | 0.14 | 0.16 | 15.9
| Analytical results for Sm in OREAS 931b (Certified Value 5.85 ppm) | [22 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  |  |  |  | 1
| No | No | 01: 11; 02: 12; 03: 13; 05: 14; 06: 17; 07: 18; 08: 20; 09: 24 |  |  |  |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 05: PF*OES/MS; 06: PF*OES/MS; 07: PF*OES; 08: PF*MS; 09: PF*MS |  |  |  |  | 2
3
| 1 | 1 | 01: 5.6122; 02: 5.9; 03: 5.43; 05: 5.4; 06: 5.2; 07: 6.05; 08: 6.1; 09: 7.7065 |  |  |  |  | 1
| 1 | 2 | 01: 5.6855; 02: 6.1; 03: 5.35; 05: 5.6; 06: 5.3; 07: 6.03; 08: 5.9; 09: 6.9986 |  |  |  |  | 17
| 1 | 3 | 01: 5.5663; 02: 5.9; 03: 5.26; 05: 6.2; 06: 5.5; 07: 5.98; 08: 6; 09: 7.6005 |  |  |  |  | 16
| 1 | 4 | 01: 6.1646; 02: 5.9; 03: 5.35; 05: 6; 06: 5; 07: 5.67; 08: 6; 09: 8.1743 |  |  |  |  | 5.8511
| 1 | 5 | 01: 5.806; 02: 5.9; 03: 5.62; 05: 6.3; 06: 5.6; 07: 6.07; 08: 6.2; 09: 7.3112 |  |  |  |  | 175
| 1 | 6 | 01: 6.1622; 02: 5.9; 03: 5.56; 05: 5.7; 06: 5.5; 07: 5.93; 08: 6; 09: 7.0039
| Mean |  | 01: 5.8328; 02: 5.9333; 03: 5.4283; 05: 5.8667; 06: 5.35; 07: 5.955; 08: 6.0333; 09: 7.4658
| Median |  | 01: 5.7458; 02: 5.9; 03: 5.39; 05: 5.85; 06: 5.4; 07: 6.005; 08: 6; 09: 7.4559
| Std Dev. |  | 01: 0.26859; 02: 0.08165; 03: 0.13761; 05: 0.3559; 06: 0.22583; 07: 0.14856; 08: 0.10328; 09: 0.45465
| Rel.Std.Dev. |  | 01: 4.605%; 02: 1.376%; 03: 2.535%; 05: 6.067%; 06: 4.221%; 07: 2.495%; 08: 1.712%; 09: 6.09%
| PDM3 |  | 01: -0.3124%; 02: 1.406%; 03: -7.225%; 05: 0.2664%; 06: -8.564%; 07: 1.776%; 08: 3.115%; 09: 27.6%
| Z-Score (Absolute) |  | 01: 0.45; 02: 0.22; 03: 3.17; 05: 0.22; 06: 3.7; 07: 0.37; 08: 0.9; 09: 10.53
| Analytical results for Sn in OREAS 931b (Certified Value 45.1 ppm) | [22 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 | [7 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [7 empty] | 1
1
| 1 | 1 | 45 | < 100 | < 100 | < 50 | 40 | 42 | 46 | < 100 | 90 | 37.4 | 45 | 57.7707 | 41 | 9.7009 | [7 empty] | 1
| 1 | 2 | 43 | < 100 | < 100 | < 50 | 49 | 44 | 46 | < 100 | 150 | 38.1 | 44 | 45.695 | 40 | 9.8703 | [7 empty] | 8
| 1 | 3 | 47 | < 100 | < 100 | 58 | 46 | 46 | 47 | < 100 | 100 | 34.8 | 43 | 52.1903 | 39 | 9.949 | [7 empty] | 16
| 1 | 4 | 47 | < 100 | < 100 | < 50 | 42 | 45 | 47 | < 100 | 100 | 40.1 | 41 | 42.1168 | 40 | 9.4312 | [7 empty] | 45.0636
| 1 | 5 | 46 | < 100 | < 100 | < 50 | 42 | 44 | 45 | < 100 | 100 | 38.2 | 43 | 39.11 | 40 | 10.612 | [7 empty] | 176
| 1 | 6 | 46 | < 100 | < 100 | < 50 | 37 | 43 | 46 | < 100 | 120 | 36.4 | 41 | 56.551 | 39 | 10.0288
| Mean |  | 45.6667 |  |  | 58 | 42.6667 | 44 | 46.1667 |  | 110 | 37.5 | 42.8333 | 48.9057 | 39.8333 | 9.932
| Median |  | 46 |  |  | 58 | 42 | 44 | 46 |  | 100 | 37.75 | 43 | 48.9427 | 40 | 9.9096
| Std Dev. |  | 1.5055 |  |  |  | 4.274 | 1.4142 | 0.75277 |  | 21.9089 | 1.7956 | 1.6021 | 7.7483 | 0.75277 | 0.39497
| Rel.Std.Dev. |  | 3.297% |  |  |  | 10.02% | 3.214% | 1.631% |  | 19.92% | 4.788% | 3.74% | 15.84% | 1.89% | 3.977%
| PDM3 |  | 1.338% |  |  | 28.71% | -5.319% | -2.36% | 2.448% |  | 144.1% | -16.78% | -4.949% | 8.526% | -11.61% | -77.96%
| Z-Score (Absolute) |  | 0.11 | 0.69 | 0.69 | 1.9 | 0.29 | 0.11 | 0.18 | 0.69 | 8.64 | 0.97 | 0.27 | 0.54 | 0.66 | 4.63
| Analytical results for Sr in OREAS 931b (Certified Value 54 ppm) | [22 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [7 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*MS | [7 empty] | 0
1
| 1 | 1 | 43 | 53.3817 | 60 | 59.3 | 39 | 40 | 54.1 | 54.3 | 53 | 56 | 62 | 62.0439 | 53 | 55.3902 | [7 empty] | 1
| 1 | 2 | 43 | 53.3289 | 59 | 59.8 | 50 | 15 | 49.9 | 52.9 | 53 | 53 | 63 | 60.1379 | 48 | 58.019 | [7 empty] | 0
| 1 | 3 | 40 | 51.0507 | 60 | 62.9 | 47 | 25 | 51.6 | 52.6 | 52 | 55 | 61 | 55.2506 | 52 | 61.5414 | [7 empty] | 16
| 1 | 4 | 44 | 50.867 | 59 | 60.4 | 50 | 20 | 52.4 | 50.5 | 50 | 55 | 60 | 53.7148 | 52 | 52.6594 | [7 empty] | 53.9512
| 1 | 5 | 43 | 53.4116 | 60 | 59.7 | 47 | 20 | 54.1 | 52.4 | 52 | 55 | 62 | 54.5478 | 51 | 58.0871 | [7 empty] | 177
| 1 | 6 | 45 | 50.9494 | 60 | 60.7 | 47 | 25 | 52.4 | 51.3 | 55 | 52 | 59 | 56.643 | 51 | 58.5296
| Mean |  | 43 | 52.1649 | 59.6667 | 60.4667 | 46.6667 | 24.1667 | 52.4167 | 52.3333 | 52.5 | 54.3333 | 61.1667 | 57.0563 | 51.1667 | 57.3711
| Median |  | 43 | 52.1898 | 60 | 60.1 | 47 | 22.5 | 52.4 | 52.5 | 52.5 | 55 | 61.5 | 55.9468 | 51.5 | 58.053
| Std Dev. |  | 1.6733 | 1.3261 | 0.5164 | 1.2941 | 4.0332 | 8.612 | 1.5917 | 1.3186 | 1.6432 | 1.5055 | 1.472 | 3.3245 | 1.7224 | 3.0262
| Rel.Std.Dev. |  | 3.891% | 2.542% | 0.8655% | 2.14% | 8.643% | 35.64% | 3.037% | 2.52% | 3.13% | 2.771% | 2.406% | 5.827% | 3.366% | 5.275%
| PDM3 |  | -20.3% | -3.311% | 10.59% | 12.08% | -13.5% | -55.21% | -2.844% | -2.999% | -2.69% | 0.7083% | 13.37% | 5.755% | -5.161% | 6.339%
| Z-Score (Absolute) |  | 1.34 | 0.04 | 1.02 | 1.14 | 0.82 | 4.02 | 0.01 | 0.02 | 0.01 | 0.27 | 1.24 | 0.65 | 0.18 | 0.7
| Analytical results for Ta in OREAS 931b (Certified Value 0.94 ppm) | [22 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [9 empty] | 2
2
| 1 | 1 | 0.99718 | 1 | < 10 | 2.9 | 0.8 | 0.8 | 0.7 | 0.8 | < 5 | 1.572 | 0.7 | 2.214 | [9 empty] | 1
| 1 | 2 | 0.76735 | 0.9 | < 10 | 4 | 0.8 | 0.9 | 0.8 | 1.7 | < 5 | 1.4867 | 0.6 | 0.94894 | [9 empty] | 10
| 1 | 3 | 0.61341 | 1.2 | 14 | 2.3 | 0.8 | 0.9 | 0.8 | 0.7 | < 5 | 1.5157 | 0.6 | 0.6145 | [9 empty] | 16
| 1 | 4 | 0.44406 | 0.9 | < 10 | 2.2 | 0.8 | 0.9 | 0.8 | 0.8 | < 5 | 1.446 | 0.6 | 0.47427 | [9 empty] | 0.94082
| 1 | 5 | 0.7164 | 1.1 | 14 | 1.4 | 0.8 | 1 | 0.8 | 1.5 | < 5 | 1.2631 | 0.6 | 1.4576 | [9 empty] | 178
| 1 | 6 | 0.81169 | 1.2 | < 10 | 1.8 | 0.8 | 0.8 | 0.8 | 0.7 | < 5 | 1.3196 | 0.6 | 3.524
| Mean |  | 0.72502 | 1.05 | 14 | 2.4333 | 0.8 | 0.88333 | 0.78333 | 1.0333 |  | 1.4339 | 0.61667 | 1.5389
| Median |  | 0.74188 | 1.05 | 14 | 2.25 | 0.8 | 0.9 | 0.8 | 0.8 |  | 1.4664 | 0.6 | 1.2033
| Std Dev. |  | 0.187 | 0.13784 | 0 | 0.91797 | 1.2162e-16 | 0.075277 | 0.040825 | 0.44572 |  | 0.1191 | 0.040825 | 1.161
| Rel.Std.Dev. |  | 25.79% | 13.13% | 0% | 37.72% | 1.52e-14% | 8.522% | 5.212% | 43.13% |  | 8.306% | 6.62% | 75.44%
| PDM3 |  | -22.94% | 11.6% | 1388% | 158.6% | -14.97% | -6.111% | -16.74% | 9.833% |  | 52.41% | -34.45% | 63.57%
| Z-Score (Absolute) |  | 0.6 | 0.02 | 13.24 | 2.65 | 0.46 | 0.3 | 0.49 | 0.02 | 2.77 | 0.75 | 0.81 | 0.95
| Analytical results for Tb in OREAS 931b (Certified Value 0.72 ppm) | [22 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab | [7 empty] | 1
| No | No | 01: 11; 02: 12; 03: 13; 10: 14; 12: 17; 13: 18; 14: 20 | [7 empty] | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 10: PF*OES/MS; 12: PF*OES/MS; 13: PF*OES; 14: PF*MS | [7 empty] | 2
3
| 1 | 1 | 01: 0.7346; 02: 0.66; 03: 0.68; 10: 0.7; 12: 0.8; 13: 0.78; 14: 0.76 | [7 empty] | 1
| 1 | 2 | 01: 0.7301; 02: 0.75; 03: 0.68; 10: 0.67; 12: 0.7; 13: 0.76; 14: 0.7 | [7 empty] | 0
| 1 | 3 | 01: 0.7225; 02: 0.71; 03: 0.66; 10: 0.71; 12: 0.7; 13: 0.75; 14: 0.73 | [7 empty] | 16
| 1 | 4 | 01: 0.7457; 02: 0.72; 03: 0.68; 10: 0.71; 12: 0.7; 13: 0.75; 14: 0.72 | [7 empty] | 0.71784
| 1 | 5 | 01: 0.7884; 02: 0.65; 03: 0.67; 10: 0.73; 12: 0.7; 13: 0.81; 14: 0.73 | [7 empty] | 179
| 1 | 6 | 01: 0.7509; 02: 0.66; 03: 0.67; 10: 0.69; 12: 0.7; 13: 0.77; 14: 0.71
| Mean |  | 01: 0.74537; 02: 0.69167; 03: 0.67333; 10: 0.70167; 12: 0.71667; 13: 0.77; 14: 0.725
| Median |  | 01: 0.74015; 02: 0.685; 03: 0.675; 10: 0.705; 12: 0.7; 13: 0.765; 14: 0.725
| Std Dev. |  | 01: 0.023471; 02: 0.040702; 03: 0.008165; 10: 0.020412; 12: 0.040825; 13: 0.022804; 14: 0.020736
| Rel.Std.Dev. |  | 01: 3.149%; 02: 5.885%; 03: 1.213%; 10: 2.909%; 12: 5.696%; 13: 2.961%; 14: 2.86%
| PDM3 |  | 01: 3.835%; 02: -3.646%; 03: -6.2%; 10: -2.253%; 12: -0.1633%; 13: 7.266%; 14: 0.9976%
| Z-Score (Absolute) |  | 01: 0.8; 02: 0.54; 03: 1; 10: 0.29; 12: NA; 13: 1.42; 14: 0.29
Results from laboratory 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 931b (Indicative Value < 1 ppm) | [22 empty] | Indicative
Te | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 10: Lab; 12: Lab; 13: Lab | [7 empty] | 1
| No | No | 01: 02; 02: 03; 03: 12; 10: 14; 12: 17; 13: 19 | [7 empty] | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*MS; 10: PF*OES/MS; 12: PF*OES/MS; 13: PF*MS | [7 empty] | 2
2
| 1 | 1 | 01: < 1; 02: < 2; 03: < 1; 10: 0.29; 12: < 6; 13: < 1 | [7 empty] | 1
| 1 | 2 | 01: < 1; 02: < 2; 03: < 1; 10: < 0.01; 12: < 6; 13: < 1 | [7 empty] | 21
| 1 | 3 | 01: < 1; 02: < 2; 03: < 1; 10: < 0.01; 12: < 6; 13: < 1 | [7 empty] | 16
| 1 | 4 | 01: < 1; 02: < 2; 03: < 1; 10: < 0.01; 12: < 6; 13: < 1 | [7 empty] | < 1
| 1 | 5 | 01: < 1; 02: < 2; 03: < 1; 10: < 0.01; 12: < 6; 13: < 1 | [7 empty] | 27
| 1 | 6 | 01: < 1; 02: < 2; 03: < 1; 10: < 0.01; 12: < 6; 13: < 1
| Mean |  | 10: 0.29
| Median |  | 10: 0.29
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | 01: 0; 02: 1.51; 03: 0; 10: 1.35; 12: 7.53; 13: 0
| Analytical results for Th in OREAS 931b (Certified Value 11.7 ppm) | [22 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [8 empty] | 1
1
| 1 | 1 | 11.2392 | 11.4 | 10.2 | 12.9132 | 11.3 | 12.4 | 12.4 | 13.3 | 11.8 | 12 | 9.7535 | 13.6 | 12.2124 | [8 empty] | 1
| 1 | 2 | 11.9097 | 11.9 | 9.94 | 12.3242 | 10.2 | 12.4 | 12.5 | 13.5 | 11.5 | 10 | 8.8188 | 12.9 | 12.9467 | [8 empty] | 11
| 1 | 3 | 11.7248 | 11.5 | 10.1 | 12.3558 | 9.9 | 12.1 | 11.6 | 13.6 | 11.9 | 11 | 11.1478 | 12.9 | 12.2923 | [8 empty] | 16
| 1 | 4 | 11.59 | 11.9 | 10.7 | 12.7122 | 9.3 | 12.4 | 11.3 | 13.5 | 11 | 11 | 9.9931 | 13.1 | 11.4576 | [8 empty] | 11.69
| 1 | 5 | 11.7992 | 12.1 | 10.3 | 13.5517 | 8.2 | 12.1 | 11.5 | 13.9 | 11.9 | 11 | 8.8053 | 12.7 | 12.0278 | [8 empty] | 180
| 1 | 6 | 11.2048 | 12.4 | 10 | 12.6447 | 10.1 | 12.2 | 11.5 | 13 | 11.6 | 13 | 10.2488 | 12.7 | 11.7692
| Mean |  | 11.578 | 11.8667 | 10.2067 | 12.7503 | 9.8333 | 12.2667 | 11.8 | 13.4667 | 11.6167 | 11.3333 | 9.7946 | 12.9833 | 12.1177
| Median |  | 11.6574 | 11.9 | 10.15 | 12.6784 | 10 | 12.3 | 11.55 | 13.5 | 11.7 | 11 | 9.8733 | 12.9 | 12.1201
| Std Dev. |  | 0.29488 | 0.37238 | 0.27471 | 0.45135 | 1.0309 | 0.15055 | 0.51381 | 0.30111 | 0.34303 | 1.0328 | 0.89548 | 0.33714 | 0.50832
| Rel.Std.Dev. |  | 2.547% | 3.138% | 2.691% | 3.54% | 10.48% | 1.227% | 4.354% | 2.236% | 2.953% | 9.113% | 9.143% | 2.597% | 4.195%
| PDM3 |  | -0.9588% | 1.511% | -12.69% | 9.07% | -15.88% | 4.933% | 0.9406% | 15.2% | -0.6277% | -3.051% | -16.21% | 11.06% | 3.658%
| Z-Score (Absolute) |  | 0.26 | 0.03 | 1.62 | 0.92 | 2 | 0.43 | 0.03 | 1.63 | 0.22 | NA | 2.04 | 1.15 | 0.28
Results from laboratory 18 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 931b (Certified Value 0.265 wt.%) | [22 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS |  | 3
3
| 1 | 1 | 0.27 | 0.27404 | 0.27 | 0.25 | 0.258 | 0.264 | 0.264 | 0.264 | 0.264 | 0.25 | 0.26 | 0.264 | 0.28 | 0.25 | 0.27 | 0.27 | 0.26 | 0.268 | 0.26 | 0.27978 |  | 1
| 1 | 2 | 0.27 | 0.28463 | 0.27 | 0.27 | 0.264 | 0.264 | 0.258 | 0.27 | 0.264 | 0.27 | 0.25 | 0.27 | 0.28 | 0.25 | 0.27 | 0.27 | 0.27 | 0.2674 | 0.25 | 0.28737 |  | 0
| 1 | 3 | 0.27 | 0.27951 | 0.27 | 0.28 | 0.264 | 0.264 | 0.264 | 0.27 | 0.264 | 0.26 | 0.25 | 0.264 | 0.28 | 0.26 | 0.27 | 0.27 | 0.27 | 0.2722 | 0.26 | 0.27495 |  | 16
| 1 | 4 | 0.28 | 0.28294 | 0.27 | 0.27 | 0.264 | 0.264 | 0.246 | 0.27 | 0.264 | 0.26 | 0.26 | 0.27 | 0.28 | 0.26 | 0.27 | 0.26 | 0.27 | 0.268 | 0.26 | 0.28323 |  | 0.26527
| 1 | 5 | 0.28 | 0.27653 | 0.26 | 0.27 | 0.264 | 0.264 | 0.264 | 0.276 | 0.27 | 0.26 | 0.25 | 0.264 | 0.29 | 0.26 | 0.27 | 0.26 | 0.27 | 0.271 | 0.26 | 0.29473 |  | 181
| 1 | 6 | 0.27 | 0.28057 | 0.27 | 0.26 | 0.264 | 0.264 | 0.264 | 0.27 | 0.264 | 0.26 | 0.25 | 0.264 | 0.28 | 0.26 | 0.27 | 0.26 | 0.27 | 0.2644 | 0.26 | 0.29749
| Mean |  | 0.27333 | 0.2797 | 0.26833 | 0.26667 | 0.263 | 0.264 | 0.26 | 0.27 | 0.265 | 0.26 | 0.25333 | 0.266 | 0.28167 | 0.25667 | 0.27 | 0.265 | 0.26833 | 0.2685 | 0.25833 | 0.28626
| Median |  | 0.27 | 0.28004 | 0.27 | 0.27 | 0.264 | 0.264 | 0.264 | 0.27 | 0.264 | 0.26 | 0.25 | 0.264 | 0.28 | 0.26 | 0.27 | 0.265 | 0.27 | 0.268 | 0.26 | 0.2853
| Std Dev. |  | 0.005164 | 0.0039425 | 0.0040825 | 0.010328 | 0.0024495 | 0 | 0.0072664 | 0.0037947 | 0.0024495 | 0.0063246 | 0.005164 | 0.0030984 | 0.0040825 | 0.005164 | 0 | 0.0054772 | 0.0040825 | 0.0027734 | 0.0040825 | 0.0086962
| Rel.Std.Dev. |  | 1.889% | 1.41% | 1.521% | 3.873% | 0.9314% | 0% | 2.795% | 1.405% | 0.9243% | 2.433% | 2.038% | 1.165% | 1.449% | 2.012% | 0% | 2.067% | 1.521% | 1.033% | 1.58% | 3.038%
| PDM3 |  | 3.041% | 5.443% | 1.156% | 0.5276% | -0.8547% | -0.4777% | -1.986% | 1.784% | -0.1007% | -1.986% | -4.499% | 0.2763% | 6.182% | -3.242% | 1.784% | -0.1007% | 1.156% | 1.219% | -2.614% | 7.915%
| Z-Score (Absolute) |  | 1.31 | 2.48 | 0.39 | 0.08 | 0.63 | 0.45 | 1.19 | 0.65 | 0.27 | 1.14 | 2.37 | 0.08 | 2.84 | 1.76 | 0.7 | 0.22 | 0.39 | 0.41 | 1.45 | 3.69
| Analytical results for Tl in OREAS 931b (Certified Value 0.74 ppm) | [22 empty] | Certified
Tl | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 04; 04: 12; 05: 14; 06: 15; 07: 17; 08: 19; 09: 20 |  | ppm
|  |  | 01: PF*MS; 02: PF*MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 08: PF*MS; 09: PF*MS |  | 2
3
| 1 | 1 | 01: 0.74075; 02: 0.8; 03: 0.85; 04: 0.6; 05: 0.7; 06: 0.7; 07: 0.7; 08: < 1; 09: 0.8 |  | 1
| 1 | 2 | 01: 0.75707; 02: 0.8; 03: 0.89; 04: 0.7; 05: 0.8; 06: 0.7; 07: 0.7; 08: < 1; 09: 0.7 |  | 0
| 1 | 3 | 01: 0.76879; 02: 0.8; 03: 0.88; 04: 0.7; 05: 0.7; 06: 0.8; 07: 0.7; 08: < 1; 09: 0.8 |  | 16
| 1 | 4 | 01: 0.7454; 02: 0.7; 03: 0.78; 04: 0.7; 05: 0.8; 06: 0.8; 07: 0.8; 08: < 1; 09: 0.8 |  | 0.74349
| 1 | 5 | 01: 0.74933; 02: 0.8; 03: 0.83; 04: 0.7; 05: 0.8; 06: 0.8; 07: 0.7; 08: < 1; 09: 0.8 |  | 182
| 1 | 6 | 01: 0.76518; 02: 0.8; 03: 0.8; 04: 0.7; 05: 0.7; 06: 0.7; 07: 0.7; 08: < 1; 09: 0.7
| Mean |  | 01: 0.75442; 02: 0.78333; 03: 0.83833; 04: 0.68333; 05: 0.75; 06: 0.75; 07: 0.71667; 09: 0.76667
| Median |  | 01: 0.7532; 02: 0.8; 03: 0.84; 04: 0.7; 05: 0.75; 06: 0.75; 07: 0.7; 09: 0.8
| Std Dev. |  | 01: 0.011166; 02: 0.040825; 03: 0.043551; 04: 0.040825; 05: 0.054772; 06: 0.054772; 07: 0.040825; 09: 0.05164
| Rel.Std.Dev. |  | 01: 1.48%; 02: 5.212%; 03: 5.195%; 04: 5.974%; 05: 7.303%; 06: 7.303%; 07: 5.696%; 09: 6.736%
| PDM3 |  | 01: 1.47%; 02: 5.359%; 03: 12.76%; 04: -8.091%; 05: 0.8758%; 06: 0.8758%; 07: -3.608%; 09: 3.117%
| Z-Score (Absolute) |  | 01: 0.09; 02: 0.67; 03: 1.79; 04: 1.35; 05: 0; 06: 0; 07: 0.67; 08: 5.06; 09: 0.34
| Analytical results for Tm in OREAS 931b (Certified Value 0.34 ppm) | [22 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  | 1
| No | No | 01: 11; 02: 12; 03: 13; 04: 14; 05: 17; 06: 18; 07: 20; 08: 24 |  | ppm
|  |  | 01: PF*OES/MS; 02: PF*MS; 03: BF*OES/MS; 04: PF*OES/MS; 05: PF*OES/MS; 06: PF*OES; 07: PF*MS; 08: PF*MS |  | 2
3
| 1 | 1 | 01: 0.3548; 02: 0.34; 03: 0.31; 04: 0.35; 05: 0.3; 06: 0.38; 07: 0.37; 08: < 2 |  | 1
| 1 | 2 | 01: 0.3756; 02: 0.32; 03: 0.3; 04: 0.35; 05: 0.3; 06: 0.37; 07: 0.35; 08: < 2 |  | 6
| 1 | 3 | 01: 0.3707; 02: 0.3; 03: 0.3; 04: 0.36; 05: 0.4; 06: 0.37; 07: 0.34; 08: < 2 |  | 16
| 1 | 4 | 01: 0.3758; 02: 0.31; 03: 0.32; 04: 0.34; 05: 0.4; 06: 0.36; 07: 0.36; 08: < 2 |  | 0.3442
| 1 | 5 | 01: 0.395; 02: 0.28; 03: 0.31; 04: 0.36; 05: 0.3; 06: 0.37; 07: 0.35; 08: < 2 |  | 183
| 1 | 6 | 01: 0.3693; 02: 0.31; 03: 0.3; 04: 0.34; 05: 0.4; 06: 0.39; 07: 0.34; 08: < 2
| Mean |  | 01: 0.37353; 02: 0.31; 03: 0.30667; 04: 0.35; 05: 0.35; 06: 0.37333; 07: 0.35167
| Median |  | 01: 0.37315; 02: 0.31; 03: 0.305; 04: 0.35; 05: 0.35; 06: 0.37; 07: 0.35
| Std Dev. |  | 01: 0.013016; 02: 0.02; 03: 0.008165; 04: 0.0089443; 05: 0.054772; 06: 0.010328; 07: 0.01169
| Rel.Std.Dev. |  | 01: 3.485%; 02: 6.452%; 03: 2.662%; 04: 2.556%; 05: 15.65%; 06: 2.766%; 07: 3.324%
| PDM3 |  | 01: 8.522%; 02: -9.936%; 03: -10.9%; 04: 1.685%; 05: 1.685%; 06: 8.464%; 07: 2.169%
| Z-Score (Absolute) |  | 01: 0.67; 02: 1.28; 03: 1.39; 04: 0.05; 05: NA; 06: 0.67; 07: 0; 08: 19.99
Results from laboratory 17 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 931b (Certified Value 3.8 ppm) | [22 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 03 | 04 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES/MS | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [8 empty] | 2
3
| 1 | 1 | 3.7645 | 5.8 | 3.62 | 3.9712 | 4 | 3.5 | 3.95 | 4.4 | 3.7 | 3 | 4.0055 | 3.83 | 4.2409 | [8 empty] | 1
| 1 | 2 | 3.8588 | 4.9 | 3.49 | 3.8397 | 4 | 3.5 | 3.9 | 4.4 | 3.7 | 4 | 3.8995 | 3.65 | 4.4175 | [8 empty] | 13
| 1 | 3 | 3.6874 | 5.2 | 3.63 | 3.7841 | 4 | 3.5 | 3.56 | 4.4 | 3.6 | 4 | 4.3048 | 3.66 | 4.2771 | [8 empty] | 16
| 1 | 4 | 3.776 | 3.3 | 3.8 | 3.8731 | 4 | 3.6 | 3.39 | 4.5 | 3.3 | 3 | 4.2687 | 3.78 | 4.2148 | [8 empty] | 3.8007
| 1 | 5 | 3.7358 | 3.3 | 3.72 | 4.1322 | 4 | 3.5 | 3.63 | 4.5 | 3.6 | 3 | 3.8335 | 3.78 | 4.4942 | [8 empty] | 184
| 1 | 6 | 3.6024 | 3.3 | 3.68 | 3.9486 | 5 | 3.5 | 3.58 | 4.4 | 3.7 | 3 | 4.0432 | 3.73 | 4.4269
| Mean |  | 3.7375 | 4.3 | 3.6567 | 3.9248 | 4.1667 | 3.5167 | 3.6683 | 4.4333 | 3.6 | 3.3333 | 4.0592 | 3.7383 | 4.3452
| Median |  | 3.7501 | 4.1 | 3.655 | 3.9108 | 4 | 3.5 | 3.605 | 4.4 | 3.65 | 3 | 4.0243 | 3.755 | 4.3473
| Std Dev. |  | 0.086808 | 1.1331 | 0.10482 | 0.12284 | 0.40825 | 0.040825 | 0.21517 | 0.05164 | 0.15492 | 0.5164 | 0.19173 | 0.071949 | 0.11543
| Rel.Std.Dev. |  | 2.323% | 26.35% | 2.866% | 3.13% | 9.798% | 1.161% | 5.866% | 1.165% | 4.303% | 15.49% | 4.723% | 1.925% | 2.657%
| PDM3 |  | -1.663% | 13.14% | -3.789% | 3.266% | 9.63% | -7.473% | -3.482% | 16.65% | -5.28% | -12.3% | 6.802% | -1.64% | 14.33%
| Z-Score (Absolute) |  | 0 | 2.03 | 0.3 | 0.67 | NA | 0.8 | 0.25 | 2.51 | 0.5 | NA | 1.16 | 0 | 2.19
Results from laboratories 12 and 18 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 931b (Certified Value 144 ppm) | [22 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | [7 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 144 | 140.9092 | 137 | 150 | 144 | 139 | 145 | 151 | 139 | 139 | 146 | 157.1911 | 141 | 188 | [7 empty] | 1
| 1 | 2 | 146 | 145.9193 | 139 | 150 | 158 | 134 | 134 | 152 | 137 | 141 | 145 | 148.7944 | 138 | 187 | [7 empty] | 0
| 1 | 3 | 146 | 147.5891 | 134 | 150 | 148 | 134 | 147 | 156 | 136 | 140 | 147 | 141.2319 | 141 | 193 | [7 empty] | 16
| 1 | 4 | 148 | 146.6146 | 141 | 150 | 152 | 139 | 143 | 155 | 139 | 144 | 141 | 143.6675 | 141 | 185 | [7 empty] | 144.0234
| 1 | 5 | 148 | 146.9026 | 144 | 150 | 151 | 134 | 143 | 158 | 139 | 143 | 148 | 138.9108 | 141 | 195 | [7 empty] | 185
| 1 | 6 | 148 | 140.1776 | 141 | 150 | 150 | 140 | 140 | 156 | 140 | 138 | 145 | 142.1556 | 140 | 189
| Mean |  | 146.6667 | 144.6854 | 139.3333 | 150 | 150.5 | 136.6667 | 142 | 154.6667 | 138.3333 | 140.8333 | 145.3333 | 145.3252 | 140.3333 | 189.5
| Median |  | 147 | 146.2669 | 140 | 150 | 150.5 | 136.5 | 143 | 155.5 | 139 | 140.5 | 145.5 | 142.9115 | 141 | 188.5
| Std Dev. |  | 1.633 | 3.261 | 3.5024 | 0 | 4.6368 | 2.9439 | 4.5607 | 2.6583 | 1.5055 | 2.3166 | 2.4221 | 6.6862 | 1.2111 | 3.7815
| Rel.Std.Dev. |  | 1.113% | 2.254% | 2.514% | 0% | 3.081% | 2.154% | 3.212% | 1.719% | 1.088% | 1.645% | 1.667% | 4.601% | 0.863% | 1.996%
| PDM3 |  | 1.835% | 0.4597% | -3.256% | 4.15% | 4.497% | -5.108% | -1.405% | 7.39% | -3.951% | -2.215% | 0.9095% | 0.9039% | -2.562% | 31.58%
| Z-Score (Absolute) |  | 0.23 | 0.04 | 0.79 | 0.7 | 0.77 | 1.16 | 0.42 | 1.35 | 0.93 | 0.58 | 0.05 | 0.04 | 0.65 | 6.21
| Analytical results for W in OREAS 931b (Certified Value 17.8 ppm) | [22 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 10 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 24 | [8 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES | PF*MS | BF*OES/MS | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | PF*MS | [8 empty] | 1
1
| 1 | 1 | 25 | 19.2024 | 17 | < 50 | 7 | 17.4 | 20 | 19 | 16.6 | 18 | 12.8563 | 20 | 17.6794 | [8 empty] | 1
| 1 | 2 | 24 | 19.2275 | 18 | < 50 | 7 | 17.6 | 20 | 19 | 16.2 | 18 | 11.9535 | 19 | 19.0285 | [8 empty] | 15
| 1 | 3 | 26 | 18.6919 | 15 | < 50 | 8 | 18.9 | 21 | 19 | 15.8 | 15 | 15.189 | 19 | 16.7877 | [8 empty] | 16
| 1 | 4 | 25 | 18.7755 | 18 | < 50 | 8 | 18.6 | 21 | 19 | 17.4 | 14 | 12.7574 | 19 | 16.7065 | [8 empty] | 17.8191
| 1 | 5 | 26 | 19.1854 | 18 | < 50 | 7 | 17.9 | 21 | 19 | 17.6 | 14 | 11.1368 | 19 | 25.1294 | [8 empty] | 186
| 1 | 6 | 25 | 18.6769 | 17 | < 50 | 10 | 18 | 20 | 19 | 16.4 | 16 | 12.6196 | 19 | 25.1433
| Mean |  | 25.1667 | 18.9599 | 17.1667 |  | 7.8333 | 18.0667 | 20.5 | 19 | 16.6667 | 15.8333 | 12.7521 | 19.1667 | 20.0791
| Median |  | 25 | 18.9805 | 17.5 |  | 7.5 | 17.95 | 20.5 | 19 | 16.5 | 15.5 | 12.6885 | 19 | 18.354
| Std Dev. |  | 0.75277 | 0.27099 | 1.169 |  | 1.169 | 0.5785 | 0.54772 | 0 | 0.70048 | 1.8348 | 1.3575 | 0.40825 | 4.0055
| Rel.Std.Dev. |  | 2.991% | 1.429% | 6.81% |  | 14.92% | 3.202% | 2.672% | 0% | 4.203% | 11.59% | 10.65% | 2.13% | 19.95%
| PDM3 |  | 41.23% | 6.402% | -3.662% |  | -56.04% | 1.389% | 15.04% | 6.627% | -6.468% | -11.14% | -28.44% | 7.562% | 12.68%
| Z-Score (Absolute) |  | 2.33 | 0 | 0.67 | 2.27 | 4.18 | 0.34 | 0.58 | 0.02 | 0.86 | 1.18 | 2.33 | 0.08 | 0.42
| Analytical results for Y in OREAS 931b (Certified Value 23.6 ppm) | [22 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 24 | [7 empty] | ppm
|  |  | PF*OES | PF*MS | PF*MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | BF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | [7 empty] | 1
2
| 1 | 1 | 23 | 24.4989 | 22.2 | 24 | 18 | 23.1218 | 18.3 | 23 | 23.7 | 24.3 | 23.1 | 25 | 20.7 | 24.9743 | [7 empty] | 1
| 1 | 2 | 23 | 24.6505 | 22.2 | 22.8 | 23 | 21.9913 | 18.9 | 22 | 22.5 | 23.6 | 23 | 25 | 20.8 | 25.6147 | [7 empty] | 16
| 1 | 3 | 23 | 23.705 | 22.6 | 23.1 | 21 | 21.996 | 18.1 | 23 | 25.3 | 23.8 | 23.8 | 24 | 20.4 | 25.1116 | [7 empty] | 16
| 1 | 4 | 23 | 22.7885 | 22.4 | 23.9 | 20 | 22.5046 | 17.3 | 24 | 23.6 | 23.1 | 22.9 | 24 | 21.1 | 25.2472 | [7 empty] | 23.5837
| 1 | 5 | 24 | 23.7851 | 23.4 | 23.9 | 21 | 23.6514 | 16.5 | 24 | 23.9 | 23.2 | 24.2 | 25 | 20.6 | 25.2254 | [7 empty] | 187
| 1 | 6 | 23 | 23.2541 | 22.7 | 23.2 | 20 | 22.8823 | 16.1 | 23 | 23.3 | 23.9 | 23.5 | 24 | 21 | 25.4201
| Mean |  | 23.1667 | 23.7804 | 22.5833 | 23.4833 | 20.5 | 22.6912 | 17.5333 | 23.1667 | 23.7167 | 23.65 | 23.4167 | 24.5 | 20.7667 | 25.2656
| Median |  | 23 | 23.7451 | 22.5 | 23.55 | 20.5 | 22.6934 | 17.7 | 23 | 23.65 | 23.7 | 23.3 | 24.5 | 20.75 | 25.2363
| Std Dev. |  | 0.40825 | 0.71273 | 0.44907 | 0.51153 | 1.6432 | 0.65603 | 1.0912 | 0.75277 | 0.91742 | 0.45056 | 0.51153 | 0.54772 | 0.2582 | 0.22639
| Rel.Std.Dev. |  | 1.762% | 2.997% | 1.989% | 2.178% | 8.015% | 2.891% | 6.223% | 3.249% | 3.868% | 1.905% | 2.184% | 2.236% | 1.243% | 0.8961%
| PDM3 |  | -1.768% | 0.834% | -4.242% | -0.4255% | -13.08% | -3.784% | -25.65% | -1.768% | 0.5639% | 0.2812% | -0.7082% | 3.885% | -11.94% | 7.132%
| Z-Score (Absolute) |  | 0.15 | 0.61 | 0.88 | 0.24 | 3.46 | 0.74 | 7.13 | 0.15 | 0.53 | 0.44 | 0.15 | 1.5 | 3.13 | 2.44
| Analytical results for Yb in OREAS 931b (Certified Value 2.25 ppm) | [22 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 11 | 12 | 13 | 14 | 17 | 18 | 20 | 24 | [13 empty] | ppm
|  |  | PF*OES/MS | PF*MS | BF*OES/MS | PF*OES/MS | PF*OES/MS | PF*OES | PF*MS | PF*MS | [13 empty] | 2
3
| 1 | 1 | 2.3631 | 2.1 | 2.14 | 2.3 | 2.2 | 2.4 | 2.3 | < 2 | [13 empty] | 1
| 1 | 2 | 2.3661 | 2.1 | 2.13 | 2.3 | 2.2 | 2.34 | 2.2 | < 2 | [13 empty] | 0
| 1 | 3 | 2.4204 | 2.1 | 2.12 | 2.2 | 2.4 | 2.36 | 2.2 | < 2 | [13 empty] | 16
| 1 | 4 | 2.4395 | 1.8 | 2.11 | 2.4 | 1.9 | 2.31 | 2.3 | < 2 | [13 empty] | 2.2463
| 1 | 5 | 2.6611 | 1.8 | 2.14 | 2.5 | 2.2 | 2.42 | 2.2 | < 2 | [13 empty] | 188
| 1 | 6 | 2.5148 | 1.9 | 2.14 | 2.4 | 2.3 | 2.37 | 2.3 | < 2
| Mean |  | 02: 2.4608; 03: 1.9667; 04: 2.13; 11: 2.35; 12: 2.2; 13: 2.3667; 14: 2.25
| Median |  | 02: 2.4299; 03: 2; 04: 2.135; 11: 2.35; 12: 2.2; 13: 2.365; 14: 2.25
| Std Dev. |  | 02: 0.11281; 03: 0.15055; 04: 0.012649; 11: 0.10488; 12: 0.16733; 13: 0.039833; 14: 0.054772
| Rel.Std.Dev. |  | 02: 4.584%; 03: 7.655%; 04: 0.5939%; 11: 4.463%; 12: 7.606%; 13: 1.683%; 14: 2.434%
| PDM3 |  | 02: 9.55%; 03: -12.45%; 04: -5.178%; 11: 4.616%; 12: -2.062%; 13: 5.358%; 14: 0.1643%
| Z-Score (Absolute) |  | 1.19 | 1.31 | 0.48 | 0.63 | 0.13 | 0.72 | 0.13 | 6.2
| Analytical results for Zn in OREAS 931b (Certified Value 283 ppm) | [22 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | 0
0
| 1 | 1 | 310 | 279.9507 | < 200000 | 300 | 330 | 280 | 290 | 250 | 280 | 239 | 274 | 285 | 283 | 300 | 290 | 300 | 227.8462 | 282 | 211 | 200 | 262.5505 | 1
| 1 | 2 | 305 | 279.4724 | < 200000 | 300 | 330 | 270 | 280 | 240 | 270 | 287 | 269 | 284 | 268 | 300 | 290 | 300 | 237.0263 | 275 | 204 | 300 | 269.0117 | 0
| 1 | 3 | 307 | 280.3775 | < 200000 | 310 | 330 | 280 | 280 | 250 | 270 | 264 | 319 | 290 | 286 | 300 | 290 | 300 | 232.8465 | 278 | 210 | 200 | 264.2983 | 16
| 1 | 4 | 309 | 287.9136 | < 200000 | 290 | 330 | 270 | 250 | 250 | 280 | 274 | 269 | 283 | 295 | 300 | 320 | 300 | 243.9623 | 287 | 215 | 300 | 267.7889 | 283.3654
| 1 | 5 | 319 | 289.4517 | < 200000 | 290 | 330 | 270 | 270 | 250 | 280 | 278 | 244 | 293 | 291 | 300 | 300 | 300 | 240.6342 | 277 | 225 | 300 | 269.6195 | 189
| 1 | 6 | 318 | 278.0229 | < 200000 | 300 | 320 | 280 | 290 | 250 | 280 | 270 | 270 | 286 | 270 | 300 | 310 | 300 | 234.332 | 285 | 218 | 200 | 269.3557
| Mean |  | 311.3333 | 282.5315 |  | 298.3333 | 328.3333 | 275 | 276.6667 | 248.3333 | 276.6667 | 268.6667 | 274.1667 | 286.8333 | 282.1667 | 300 | 300 | 300 | 236.1079 | 280.6667 | 213.8333 | 250 | 267.1041
| Median |  | 309.5 | 280.1641 |  | 300 | 330 | 275 | 280 | 250 | 280 | 272 | 269.5 | 285.5 | 284.5 | 300 | 295 | 300 | 235.6791 | 280 | 213 | 250 | 268.4003
| Std Dev. |  | 5.8195 | 4.8547 |  | 7.5277 | 4.0825 | 5.4772 | 15.0555 | 4.0825 | 5.164 | 16.4641 | 24.4574 | 3.8687 | 11.0167 | 0 | 12.6491 | 0 | 5.7457 | 4.761 | 7.2503 | 54.7723 | 2.9703
| Rel.Std.Dev. |  | 1.869% | 1.718% |  | 2.523% | 1.243% | 1.992% | 5.442% | 1.644% | 1.866% | 6.128% | 8.921% | 1.349% | 3.904% | 0% | 4.216% | 0% | 2.434% | 1.696% | 3.391% | 21.91% | 1.112%
| PDM3 |  | 9.87% | -0.2943% |  | 5.282% | 15.87% | -2.952% | -2.364% | -12.36% | -2.364% | -5.187% | -3.246% | 1.224% | -0.4231% | 5.87% | 5.87% | 5.87% | -16.68% | -0.9524% | -24.54% | -11.77% | -5.739%
| Z-Score (Absolute) |  | 1.17 | 0.07 | 3806.13 | 0.67 | 1.82 | 0.22 | 0.15 | 1.23 | 0.15 | 0.46 | 0.25 | 0.24 | 0.06 | 0.74 | 0.74 | 0.74 | 1.7 | 0 | 2.55 | 1.17 | 0.52
| Analytical results for Zr in OREAS 931b (Indicative Value 82 ppm) | [22 empty] | Indicative
Zr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab | 1
| No | No | 01: 13; 02: 14; 03: 20; 04: 24 | ppm
|  |  | 01: BF*OES/MS; 02: PF*OES/MS; 03: PF*MS; 04: PF*MS | 0
0
| 1 | 1 | 01: 94.4; 02: 88.6; 03: 77.6; 04: 67.4765 | 1
| 1 | 2 | 01: 91.1; 02: 88.3; 03: 76.4; 04: 66.9493 | 19
| 1 | 3 | 01: 96.3; 02: 92.1; 03: 74.3; 04: 65.8282 | 16
| 1 | 4 | 01: 94.5; 02: 89.2; 03: 76.1; 04: 64.06 | 82.195
| 1 | 5 | 01: 89.2; 02: 104; 03: 76; 04: 79.3161 | 25
| 1 | 6 | 01: 90.5; 02: 90.7; 03: 73.7; 04: 80.2706
| Mean |  | 01: 92.6667; 02: 92.15; 03: 75.6833; 04: 70.6501
| Median |  | 01: 92.75; 02: 89.95; 03: 76.05; 04: 67.2129
| Std Dev. |  | 01: 2.7833; 02: 5.9775; 03: 1.4359; 04: 7.1847
| Rel.Std.Dev. |  | 01: 3.004%; 02: 6.487%; 03: 1.897%; 04: 10.17%
| PDM3 |  | 01: 12.74%; 02: 12.11%; 03: -7.922%; 04: -14.05%
| Z-Score (Absolute) |  | 01: 0.69; 02: 0.65; 03: 0.65; 04: 1.05
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C253: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C271: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C289: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C325: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C343: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C380: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C434: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C452: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C470: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C488: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C506: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C524: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C542: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C560: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C578: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C596: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C614: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C632: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C650: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C668: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C687: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C705: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C723: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C741: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C759: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C778: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C796: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C814: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C832: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C850: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C868: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C886: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C904: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C923: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C941: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C960: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C978: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C996: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1015: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1034: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1052: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1070: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1088: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1106: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1124: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 931b (Indicative Value 5.21 wt.%) | [5 empty] | Indicative
LOI1000 | Round | Replicate | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 13 | 14 | 19 | wt.%
|  |  | BF*XRF | BF*OES | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 5.3 | 5.17 | 5.32 | 5.0164 | 1
| 1 | 2 | 5.4 | 5.14 | 5.32 | 5.1491 | 23
| 1 | 3 | 5.35 | 4.99 | 5.36 | 5.2202 | 16
| 1 | 4 | 5.31 | 4.99 | 5.36 | 4.9143 | 5.2115
| 1 | 5 | 5.43 | 5.07 | 5.27 | 5.0442 | 29
| 1 | 6 | 5.37 | 5.15 | 5.28 | 5.1509
| Mean |  | 5.36 | 5.085 | 5.3183 | 5.0825
| Median |  | 5.36 | 5.105 | 5.32 | 5.0966
| Std Dev. |  | 0.050596 | 0.080932 | 0.038166 | 0.11153
| Rel.Std.Dev. |  | 0.944% | 1.592% | 0.7176% | 2.194%
| PDM3 |  | 2.85% | -2.427% | 2.051% | -2.474%
| Z-Score (Absolute) |  | 0.91 | 0.67 | 0.67 | 0.68
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 931b (Indicative Value 0.679 wt.%) | [23 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | Lab | [18 empty] | 1
| No | No | 13 | 15 | 19 | 24 | [18 empty] | wt.%
|  |  | IRC | IRC | IRC | IRC | [18 empty] | 3
3
| 1 | 1 | 0.68 | 0.69 | 0.668 | 0.66469 | [18 empty] | 1
| 1 | 2 | 0.68 | 0.67 | 0.674 | 0.68044 | [18 empty] | 25
| 1 | 3 | 0.67 | 0.68 | 0.681 | 0.68634 | [18 empty] | 16
| 1 | 4 | 0.67 | 0.68 | 0.669 | 0.69584 | [18 empty] | 0.67919
| 1 | 5 | 0.7 | 0.67 | 0.675 | 0.66681 | [18 empty] | 31
| 1 | 6 | 0.69 | 0.69 | 0.682 | 0.6875
| Mean |  | 0.68167 | 0.68 | 0.67483 | 0.68027
| Median |  | 0.68 | 0.68 | 0.6745 | 0.68339
| Std Dev. |  | 0.01169 | 0.0089443 | 0.0058452 | 0.012293
| Rel.Std.Dev. |  | 1.715% | 1.315% | 0.8662% | 1.807%
| PDM3 |  | 0.3643% | 0.119% | -0.6418% | 0.1585%
| Z-Score (Absolute) |  | 1.24 | 0.11 | 4.29 | 0.11
| Analytical results for S in OREAS 931b (Certified Value 4.43 wt.%) | [23 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 24 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 4.42 | 4.6509 | 4.31 | 4.4 | 4.38 | 4.33 | 4.52 | 4.44 | 4.33 | 4.12 | 3.902 | 4.43 | 4.4 | 4.69 | 4.44 | 4.13 | 4.48 | 4.6 | 4.405 | 4.527 | 4.24 | 4.4158 | 1
| 1 | 2 | 4.42 | 4.6526 | 4.24 | 4.44 | 4.4 | 4.36 | 4.55 | 4.52 | 4.32 | 4.14 | 3.774 | 4.45 | 4.35 | 4.71 | 4.44 | 4.52 | 4.52 | 4.65 | 4.42 | 4.437 | 4.23 | 4.1981 | 0
| 1 | 3 | 4.42 | 4.6357 | 4.25 | 4.45 | 4.32 | 4.46 | 4.54 | 4.35 | 4.28 | 4.04 | 3.945 | 4.45 | 4.34 | 4.7 | 4.47 | 4.52 | 4.42 | 4.58 | 4.375 | 4.465 | 4.24 | 4.4585 | 16
| 1 | 4 | 4.45 | 4.6403 | 4.27 | 4.42 | 4.44 | 4.6 | 4.54 | 4.54 | 4.33 | 4.18 | 3.922 | 4.44 | 4.26 | 4.71 | 4.43 | 4.41 | 4.47 | 4.6 | 4.45 | 4.453 | 4.24 | 4.6246 | 4.4299
| 1 | 5 | 4.46 | 4.6316 | 4.28 | 4.39 | 4.29 | 4.49 | 4.49 | 4.51 | 4.36 | 4.26 | 3.902 | 4.44 | 4.3 | 4.74 | 4.44 | 4.33 | 4.53 | 4.61 | 4.515 | 4.441 | 4.25 | 4.0195 | 191
| 1 | 6 | 4.44 | 4.6619 | 4.32 | 4.42 | 4.36 | 4.35 | 4.49 | 4.47 | 4.35 | 4.21 | 3.895 | 4.45 | 4.31 | 4.73 | 4.45 | 4.48 | 4.44 | 4.61 | 4.515 | 4.398 | 4.27 | 4.4424
| Mean |  | 4.435 | 4.6455 | 4.2783 | 4.42 | 4.365 | 4.4317 | 4.5217 | 4.4717 | 4.3283 | 4.1583 | 3.89 | 4.4433 | 4.3267 | 4.7133 | 4.445 | 4.3983 | 4.4767 | 4.6083 | 4.4467 | 4.4535 | 4.245 | 4.3598
| Median |  | 4.43 | 4.6456 | 4.275 | 4.42 | 4.37 | 4.41 | 4.53 | 4.49 | 4.33 | 4.16 | 3.902 | 4.445 | 4.325 | 4.71 | 4.44 | 4.445 | 4.475 | 4.605 | 4.435 | 4.447 | 4.24 | 4.4291
| Std Dev. |  | 0.017607 | 0.011536 | 0.031885 | 0.022804 | 0.054314 | 0.10458 | 0.026394 | 0.06969 | 0.027869 | 0.076529 | 0.059696 | 0.008165 | 0.048028 | 0.018619 | 0.013784 | 0.15039 | 0.043205 | 0.023166 | 0.058195 | 0.042529 | 0.013784 | 0.2153
| Rel.Std.Dev. |  | 0.397% | 0.2483% | 0.7453% | 0.5159% | 1.244% | 2.36% | 0.5837% | 1.558% | 0.6439% | 1.84% | 1.535% | 0.1838% | 1.11% | 0.395% | 0.3101% | 3.419% | 0.9651% | 0.5027% | 1.309% | 0.955% | 0.3247% | 4.938%
| PDM3 |  | 0.1146% | 4.866% | -3.422% | -0.224% | -1.466% | 0.03939% | 2.071% | 0.9423% | -2.293% | -6.131% | -12.19% | 0.3028% | -2.331% | 6.398% | 0.3404% | -0.7131% | 1.055% | 4.027% | 0.378% | 0.5323% | -4.174% | -1.583%
| Z-Score (Absolute) |  | 0.02 | 2.02 | 1.47 | 0.13 | 0.65 | 0.02 | 0.84 | 0.36 | 1 | 2.61 | 5.17 | 0.1 | 1.01 | 2.66 | 0.11 | 0.33 | 0.41 | 1.66 | 0.13 | 0.19 | 1.79 | 0.7
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.