## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 241b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 7.1298 | 7.0665 | 7.1931 | 7.1065 | 7.1531
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 6.9603 | 6.8674 | 7.0532 | 6.9354 | 6.9852
Cyanide Leach
Au, Gold (ppm) | 6.8738 | 6.7988 | 6.9488 | 6.8651 | 6.8825
4-Acid Digestion
Ag, Silver (ppm) | 1.7225 | 1.618 | 1.827 | 1.6683 | 1.7767
Al, Aluminium (wt.%) | 6.5128 | 6.322 | 6.7036 | 6.382 | 6.6436
As, Arsenic (ppm) | 86.8202 | 83.5184 | 90.1221 | 84.1739 | 89.4665
Ba, Barium (ppm) | 256.9674 | 249.1295 | 264.8053 | 251.9328 | 262.002
Be, Beryllium (ppm) | 0.4644 | 0.42551 | 0.50328 | 0.43792 | 0.49087
Bi, Bismuth (ppm) | 0.069047 | 0.058785 | 0.079309 | IND | IND
Ca, Calcium (wt.%) | 5.7858 | 5.6058 | 5.9658 | 5.6403 | 5.9314
Cd, Cadmium (ppm) | 0.77953 | 0.73367 | 0.82538 | 0.74838 | 0.81068
Ce, Cerium (ppm) | 14.8499 | 14.1364 | 15.5634 | 14.4304 | 15.2694
Co, Cobalt (ppm) | 39.7593 | 38.1834 | 41.3353 | 38.5227 | 40.996
Cr, Chromium (ppm) | 119.3824 | 113.3486 | 125.4163 | 114.8738 | 123.8911
Cs, Caesium (ppm) | 1.2178 | 1.1587 | 1.2769 | 1.181 | 1.2547
Cu, Copper (ppm) | 166.3441 | 161.3701 | 171.3182 | 162.0659 | 170.6224
Dy, Dysprosium (ppm) | 3.6304 | 3.3984 | 3.8623 | 3.5229 | 3.7378
Er, Erbium (ppm) | 2.209 | 2.0141 | 2.4039 | 2.1177 | 2.3003
Eu, Europium (ppm) | 0.9703 | 0.86482 | 1.0758 | 0.9152 | 1.0254
Fe, Iron (wt.%) | 7.5421 | 7.331 | 7.7532 | 7.4065 | 7.6777
Ga, Gallium (ppm) | 15.334 | 14.5408 | 16.1272 | 14.8146 | 15.8534
Gd, Gadolinium (ppm) | 3.058 | 2.8429 | 3.2731 | 2.9288 | 3.1872
Ge, Germanium (ppm) | 0.10253 | 0.07025 | 0.13482 | IND | IND
Hf, Hafnium (ppm) | 1.7504 | 1.6262 | 1.8747 | 1.6747 | 1.8261
Ho, Holmium (ppm) | 0.75788 | 0.69576 | 0.82 | 0.7255 | 0.79026
In, Indium (ppm) | 0.078107 | 0.07118 | 0.085034 | 0.071329 | 0.084886
K, Potassium (wt.%) | 0.62862 | 0.60926 | 0.64799 | 0.6159 | 0.64134
La, Lanthanum (ppm) | 6.3353 | 6.0181 | 6.6525 | 6.1845 | 6.4861
Li, Lithium (ppm) | 12.8257 | 12.3645 | 13.287 | 12.3375 | 13.314
Lu, Lutetium (ppm) | 0.31966 | 0.28473 | 0.35458 | 0.30851 | 0.3308
Mg, Magnesium (wt.%) | 3.3075 | 3.2041 | 3.4109 | 3.2337 | 3.3812
Mn, Manganese (wt.%) | 0.13298 | 0.1288 | 0.13716 | 0.12985 | 0.13611
Mo, Molybdenum (ppm) | 2.193 | 2.086 | 2.2999 | 2.0433 | 2.3426
Na, Sodium (wt.%) | 1.9265 | 1.8652 | 1.9877 | 1.862 | 1.9909
Nb, Niobium (ppm) | 3.6613 | 3.4707 | 3.8518 | 3.4638 | 3.8587
Nd, Neodymium (ppm) | 9.2007 | 8.3988 | 10.0026 | 8.9279 | 9.4736
Ni, Nickel (ppm) | 77.4656 | 75.0027 | 79.9285 | 75.3393 | 79.5919
P, Phosphorus (wt.%) | 0.044165 | 0.042884 | 0.045447 | 0.043143 | 0.045187
Pb, Lead (ppm) | 37.3356 | 36.2188 | 38.4524 | 36.1516 | 38.5196
Pr, Praseodymium (ppm) | 1.9902 | 1.8018 | 2.1786 | 1.9085 | 2.0719
Rb, Rubidium (ppm) | 18.0996 | 17.481 | 18.7182 | 17.4941 | 18.7052
Re, Rhenium (ppm) | 0.0025963 | 0.0017758 | 0.0034168 | IND | IND
S, Sulphur (wt.%) | 0.49638 | 0.47873 | 0.51403 | 0.4848 | 0.50796
Sb, Antimony (ppm) | 1.9844 | 1.8853 | 2.0835 | 1.8594 | 2.1093
Sc, Scandium (ppm) | 35.2829 | 34.1452 | 36.4205 | 34.2708 | 36.2949
Sm, Samarium (ppm) | 2.6059 | 2.4398 | 2.7721 | 2.466 | 2.7459
Sn, Tin (ppm) | 1.3151 | 1.1913 | 1.4388 | 1.1855 | 1.4446
Sr, Strontium (ppm) | 155.6961 | 150.7039 | 160.6882 | 151.3627 | 160.0294
Ta, Tantalum (ppm) | 0.25106 | 0.23155 | 0.27057 | 0.23502 | 0.26711
Tb, Terbium (ppm) | 0.53307 | 0.49649 | 0.56966 | 0.51456 | 0.55159
Te, Tellurium (ppm) | 0.13309 | 0.095634 | 0.17055 | IND | IND
Th, Thorium (ppm) | 1.3048 | 1.2251 | 1.3844 | 1.2672 | 1.3424
Ti, Titanium (wt.%) | 0.59234 | 0.57575 | 0.60892 | 0.57776 | 0.60691
Tl, Thallium (ppm) | 0.30994 | 0.29349 | 0.3264 | 0.29832 | 0.32157
Tm, Thulium (ppm) | 0.31763 | 0.28468 | 0.35058 | 0.29851 | 0.33675
U, Uranium (ppm) | 0.39681 | 0.38768 | 0.40595 | 0.38023 | 0.41339
V, Vanadium (ppm) | 250.4634 | 243.9302 | 256.9967 | 244.886 | 256.0409
W, Tungsten (ppm) | 38.2764 | 36.5139 | 40.039 | 37.1808 | 39.3721
Y, Yttrium (ppm) | 19.5048 | 18.6737 | 20.3359 | 18.997 | 20.0126
Yb, Ytterbium (ppm) | 2.1017 | 1.9614 | 2.2419 | 2.0223 | 2.181
Zn, Zinc (ppm) | 162.9386 | 158.6458 | 167.2315 | 158.8371 | 167.0401
Zr, Zirconium (ppm) | 57.3284 | 53.9282 | 60.7286 | 55.1975 | 59.4592
Aqua Regia Digestion
Ag, Silver (ppm) | 1.7333 | 1.6668 | 1.7999 | 1.6679 | 1.7988
Al, Aluminium (wt.%) | 3.044 | 2.9275 | 3.1606 | 2.9805 | 3.1075
As, Arsenic (ppm) | 85.5163 | 82.4108 | 88.6218 | 83.4433 | 87.5894
B, Boron (ppm) | 29.9367 | 25.0089 | 34.8646 | 27.8479 | 32.0256
Ba, Barium (ppm) | 40.7607 | 39.5317 | 41.9898 | 39.644 | 41.8775
Be, Beryllium (ppm) | 0.25146 | 0.224 | 0.27892 | 0.23868 | 0.26423
Bi, Bismuth (ppm) | 0.062926 | 0.054457 | 0.071395 | IND | IND
Ca, Calcium (wt.%) | 2.3084 | 2.188 | 2.4288 | 2.2547 | 2.3622
Cd, Cadmium (ppm) | 0.77943 | 0.73647 | 0.82238 | 0.75701 | 0.80185
Ce, Cerium (ppm) | 11.2415 | 10.8486 | 11.6344 | 10.9592 | 11.5239
Co, Cobalt (ppm) | 28.7241 | 27.6186 | 29.8296 | 27.9734 | 29.4748
Cr, Chromium (ppm) | 33.7898 | 32.1091 | 35.4705 | 32.4864 | 35.0932
Cs, Caesium (ppm) | 0.8851 | 0.83557 | 0.93464 | 0.86078 | 0.90943
Cu, Copper (ppm) | 165.3704 | 161.0604 | 169.6805 | 161.7584 | 168.9825
Dy, Dysprosium (ppm) | 2.1191 | 1.7564 | 2.4818 | 2.0199 | 2.2184
Er, Erbium (ppm) | 1.2188 | 1.0068 | 1.4308 | 1.1825 | 1.2551
Eu, Europium (ppm) | 0.48373 | 0.39614 | 0.57132 | 0.4653 | 0.50216
Fe, Iron (wt.%) | 5.5679 | 5.4426 | 5.6932 | 5.4663 | 5.6695
Ga, Gallium (ppm) | 10.0383 | 9.4538 | 10.6228 | 9.7897 | 10.2869
Gd, Gadolinium (ppm) | 1.9099 | 1.6459 | 2.1739 | 1.8294 | 1.9904
Ge, Germanium (ppm) | 0.1019 | 0.078165 | 0.12563 | IND | IND
Hf, Hafnium (ppm) | 0.45457 | 0.41083 | 0.49832 | 0.43475 | 0.4744
Ho, Holmium (ppm) | 0.4233 | 0.32866 | 0.51794 | 0.4008 | 0.4458
In, Indium (ppm) | 0.038595 | 0.034596 | 0.042593 | 0.034084 | 0.043105
K, Potassium (wt.%) | 0.18316 | 0.17597 | 0.19036 | 0.1786 | 0.18772
La, Lanthanum (ppm) | 5.1149 | 4.8932 | 5.3366 | 4.9678 | 5.2621
Li, Lithium (ppm) | 9.8497 | 9.3451 | 10.3543 | 9.5597 | 10.1396
Lu, Lutetium (ppm) | 0.14624 | 0.11072 | 0.18176 | IND | IND
Mg, Magnesium (wt.%) | 1.5763 | 1.5279 | 1.6248 | 1.5423 | 1.6104
Mn, Manganese (wt.%) | 0.07382 | 0.071354 | 0.076285 | 0.072476 | 0.075163
Mo, Molybdenum (ppm) | 2.1628 | 2.0752 | 2.2504 | 2.0679 | 2.2576
Na, Sodium (wt.%) | 0.28305 | 0.27395 | 0.29215 | 0.27072 | 0.29538
Nd, Neodymium (ppm) | 6.5174 | 5.9923 | 7.0426 | 6.3603 | 6.6746
Ni, Nickel (ppm) | 54.7118 | 52.6541 | 56.7695 | 53.2871 | 56.1365
P, Phosphorus (wt.%) | 0.042972 | 0.041721 | 0.044222 | 0.041982 | 0.043961
Pb, Lead (ppm) | 37.4093 | 36.1667 | 38.6518 | 36.4684 | 38.3501
Pr, Praseodymium (ppm) | 1.4185 | 1.2661 | 1.5709 | 1.3779 | 1.4591
Rb, Rubidium (ppm) | 7.9145 | 7.5797 | 8.2493 | 7.7275 | 8.1015
Re, Rhenium (ppm) | 0.00204 | 0.0012334 | 0.0028466 | IND | IND
S, Sulphur (wt.%) | 0.49617 | 0.48313 | 0.5092 | 0.48525 | 0.50708
Sb, Antimony (ppm) | 1.2173 | 1.0466 | 1.3881 | 1.1242 | 1.3105
Sc, Scandium (ppm) | 5.3109 | 4.938 | 5.6839 | 5.149 | 5.4728
Sm, Samarium (ppm) | 1.5959 | 1.378 | 1.8138 | 1.5563 | 1.6355
Sn, Tin (ppm) | 0.73396 | 0.65764 | 0.81027 | 0.68989 | 0.77802
Sr, Strontium (ppm) | 37.4521 | 34.9588 | 39.9453 | 35.8055 | 39.0987
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.31475 | 0.29714 | 0.33236 | 0.30329 | 0.32621
Te, Tellurium (ppm) | 0.12598 | 0.097841 | 0.15412 | IND | IND
Th, Thorium (ppm) | 1.0542 | 0.97318 | 1.1353 | 1.0125 | 1.0959
Ti, Titanium (wt.%) | 0.30233 | 0.27825 | 0.32641 | 0.29122 | 0.31344
Tl, Thallium (ppm) | 0.16316 | 0.15349 | 0.17282 | 0.13011 | 0.1962
Tm, Thulium (ppm) | 0.16728 | 0.14002 | 0.19454 | IND | IND
U, Uranium (ppm) | 0.25982 | 0.24316 | 0.27647 | 0.24336 | 0.27627
V, Vanadium (ppm) | 131.1176 | 123.6012 | 138.6341 | 128.1196 | 134.1157
W, Tungsten (ppm) | 27.1493 | 24.9748 | 29.3238 | 26.329 | 27.9696
Y, Yttrium (ppm) | 10.9312 | 10.5044 | 11.358 | 10.7006 | 11.1618
Yb, Ytterbium (ppm) | 1.0321 | 0.89629 | 1.168 | 0.98166 | 1.0826
Zn, Zinc (ppm) | 149.3803 | 144.1234 | 154.6372 | 146.1834 | 152.5772
Zr, Zirconium (ppm) | 15.9799 | 14.6015 | 17.3583 | 15.3807 | 16.5792
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 241b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 11.4675 | Pt | ppb | 9.3569
4-Acid Digestion
B | ppm | 35.6707 | Hg | ppm | 0.076987 | Se | ppm | 0.80073
Aqua Regia Digestion
Hg | ppm | 0.056923 | Pd | ppb | 10.5917 | Si | wt.% | 0.11038
Ir | ppm | < 0.003 | Pt | ppb | 8.5458
Nb | ppm | 0.12471 | Se | ppm | 0.7056
Borate Fusion XRF
Al2O3 | wt.% | 12.39 | MgO | wt.% | 5.615 | S | wt.% | 0.41246
CaO | wt.% | 8.36 | MnO | wt.% | 0.1615 | SiO2 | wt.% | 54.395
Fe2O3 | wt.% | 11.155 | Na2O | wt.% | 2.59 | TiO2 | wt.% | 1.04
K2O | wt.% | 0.7345 | P2O5 | wt.% | 0.0965
Thermogravimetry
LOI1000 | wt.% | 3.235
Infrared Combustion
C | wt.% | 0.17 | S | wt.% | 0.465
Laser Ablation ICP-MS
Ag | ppm | 1.8 | Hf | ppm | 2.145 | Sm | ppm | 2.66
As | ppm | 85.3 | Ho | ppm | 0.825 | Sn | ppm | 1.3
Ba | ppm | 249.5 | In | ppm | 0.075 | Sr | ppm | 152
Be | ppm | 0.6 | La | ppm | 6.68 | Ta | ppm | 0.24
Bi | ppm | 0.06 | Lu | ppm | 0.325 | Tb | ppm | 0.55
Cd | ppm | 0.8 | Mn | wt.% | 0.1405 | Te | ppm | < 0.2
Ce | ppm | 14.3 | Mo | ppm | 2.2 | Th | ppm | 1.31
Co | ppm | 41.65 | Nb | ppm | 3.79 | Ti | wt.% | 0.618
Cr | ppm | 136 | Nd | ppm | 9.315 | Tl | ppm | 0.2
Cs | ppm | 1.165 | Ni | ppm | 87 | Tm | ppm | 0.34
Cu | ppm | 169 | Pb | ppm | 38 | U | ppm | 0.37
Dy | ppm | 3.76 | Pr | ppm | 2.02 | V | ppm | 262.5
Er | ppm | 2.405 | Rb | ppm | 17.6 | W | ppm | 38
Eu | ppm | 0.97 | Re | ppm | 0.015 | Y | ppm | 20.6
Ga | ppm | 14.7 | Sb | ppm | 2.1 | Yb | ppm | 2.355
Gd | ppm | 3.285 | Sc | ppm | 36.35 | Zn | ppm | 172.5
Ge | ppm | 1.55 | Se | ppm | < 5 | Zr | ppm | 74.5
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 241b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 7.1298 | 0.18156 | 6.7667 | 7.4929 | 6.5851 | 7.6745 | 2.546% | 5.093% | 7.639% | 6.7733 | 7.4863
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 6.9603 | 0.21943 | 6.5214 | 7.3991 | 6.302 | 7.6186 | 3.153% | 6.305% | 9.458% | 6.6123 | 7.3083
Cyanide Leach
Au, ppm | 6.8738 | 0.16046 | 6.5529 | 7.1947 | 6.3924 | 7.3552 | 2.334% | 4.669% | 7.003% | 6.5301 | 7.2175
4-Acid Digestion
Ag, ppm | 1.7225 | 0.088887 | 1.5447 | 1.9003 | 1.4558 | 1.9892 | 5.16% | 10.32% | 15.48% | 1.6364 | 1.8086
Al, wt.% | 6.5128 | 0.23461 | 6.0436 | 6.982 | 5.809 | 7.2166 | 3.602% | 7.205% | 10.81% | 6.1872 | 6.8384
As, ppm | 86.8202 | 3.2936 | 80.2331 | 93.4074 | 76.9395 | 96.701 | 3.794% | 7.587% | 11.38% | 82.4792 | 91.1613
Ba, ppm | 256.9674 | 9.394 | 238.1795 | 275.7553 | 228.7855 | 285.1493 | 3.656% | 7.311% | 10.97% | 244.119 | 269.8158
Be, ppm | 0.4644 | 0.05258 | 0.35923 | 0.56956 | 0.30665 | 0.62214 | 11.32% | 22.64% | 33.97% | 0.44118 | 0.48761
Bi, ppm | 0.069047 | 0.010425 | 0.048197 | 0.089897 | 0.037772 | 0.10032 | 15.1% | 30.2% | 45.29% | 0.065595 | 0.0725
Ca, wt.% | 5.7858 | 0.26643 | 5.253 | 6.3187 | 4.9865 | 6.5851 | 4.605% | 9.21% | 13.81% | 5.4965 | 6.0751
Cd, ppm | 0.77953 | 0.035358 | 0.70881 | 0.85024 | 0.67345 | 0.8856 | 4.536% | 9.072% | 13.61% | 0.74055 | 0.8185
Ce, ppm | 14.8499 | 1.0924 | 12.6651 | 17.0347 | 11.5726 | 18.1272 | 7.356% | 14.71% | 22.07% | 14.1074 | 15.5924
Co, ppm | 39.7593 | 2.2644 | 35.2304 | 44.2882 | 32.966 | 46.5527 | 5.695% | 11.39% | 17.09% | 37.7714 | 41.7473
Cr, ppm | 119.3824 | 11.9719 | 95.4387 | 143.3262 | 83.4669 | 155.298 | 10.03% | 20.06% | 30.08% | 113.4133 | 125.3516
Cs, ppm | 1.2178 | 0.077485 | 1.0629 | 1.3728 | 0.98538 | 1.4503 | 6.363% | 12.73% | 19.09% | 1.1569 | 1.2787
Cu, ppm | 166.3441 | 7.3828 | 151.5785 | 181.1098 | 144.1956 | 188.4927 | 4.438% | 8.877% | 13.31% | 158.0269 | 174.6613
Dy, ppm | 3.6304 | 0.26981 | 3.0908 | 4.17 | 2.8209 | 4.4398 | 7.432% | 14.86% | 22.3% | 3.4489 | 3.8119
Er, ppm | 2.209 | 0.20415 | 1.8007 | 2.6173 | 1.5966 | 2.8215 | 9.242% | 18.48% | 27.73% | 2.0986 | 2.3195
Eu, ppm | 0.9703 | 0.11901 | 0.73228 | 1.2083 | 0.61327 | 1.3273 | 12.27% | 24.53% | 36.8% | 0.92179 | 1.0188
Fe, wt.% | 7.5421 | 0.28401 | 6.9741 | 8.1101 | 6.6901 | 8.3942 | 3.766% | 7.531% | 11.3% | 7.165 | 7.9192
Ga, ppm | 15.334 | 0.84235 | 13.6493 | 17.0187 | 12.807 | 17.8611 | 5.493% | 10.99% | 16.48% | 14.5673 | 16.1007
Gd, ppm | 3.058 | 0.20661 | 2.6448 | 3.4713 | 2.4382 | 3.6779 | 6.756% | 13.51% | 20.27% | 2.9051 | 3.2109
Ge, ppm | 0.10253 | 0.036271 | 0.029993 | 0.17508 | 0 | 0.21135 | 35.37% | 70.75% | 106.1% | 0.097407 | 0.10766
Hf, ppm | 1.7504 | 0.13169 | 1.487 | 2.0138 | 1.3553 | 2.1455 | 7.523% | 15.05% | 22.57% | 1.6629 | 1.8379
Ho, ppm | 0.75788 | 0.06061 | 0.63666 | 0.8791 | 0.57605 | 0.93971 | 7.997% | 15.99% | 23.99% | 0.71999 | 0.79578
In, ppm | 0.078107 | 0.0061983 | 0.065711 | 0.090504 | 0.059512 | 0.096702 | 7.936% | 15.87% | 23.81% | 0.074202 | 0.082013
K, wt.% | 0.62862 | 0.028636 | 0.57135 | 0.68589 | 0.54271 | 0.71453 | 4.555% | 9.111% | 13.67% | 0.59719 | 0.66005
La, ppm | 6.3353 | 0.39131 | 5.5527 | 7.1179 | 5.1614 | 7.5092 | 6.177% | 12.35% | 18.53% | 6.0185 | 6.6521
Li, ppm | 12.8257 | 0.55171 | 11.7223 | 13.9292 | 11.1706 | 14.4809 | 4.302% | 8.603% | 12.9% | 12.1845 | 13.467
Lu, ppm | 0.31966 | 0.028627 | 0.2624 | 0.37691 | 0.23378 | 0.40554 | 8.955% | 17.91% | 26.87% | 0.30367 | 0.33564
Mg, wt.% | 3.3075 | 0.12342 | 3.0606 | 3.5543 | 2.9372 | 3.6777 | 3.732% | 7.463% | 11.19% | 3.1421 | 3.4728
Mn, wt.% | 0.13298 | 0.0041814 | 0.12462 | 0.14134 | 0.12044 | 0.14552 | 3.144% | 6.289% | 9.433% | 0.12633 | 0.13963
Mo, ppm | 2.193 | 0.15289 | 1.8872 | 2.4987 | 1.7343 | 2.6516 | 6.972% | 13.94% | 20.92% | 2.0833 | 2.3026
Na, wt.% | 1.9265 | 0.077963 | 1.7705 | 2.0824 | 1.6926 | 2.1604 | 4.047% | 8.094% | 12.14% | 1.8301 | 2.0228
Nb, ppm | 3.6613 | 0.18405 | 3.2931 | 4.0294 | 3.1091 | 4.2134 | 5.027% | 10.05% | 15.08% | 3.4782 | 3.8443
Nd, ppm | 9.2007 | 0.7265 | 7.7477 | 10.6537 | 7.0212 | 11.3802 | 7.896% | 15.79% | 23.69% | 8.7407 | 9.6608
Ni, ppm | 77.4656 | 2.2652 | 72.9353 | 81.9959 | 70.6701 | 84.2611 | 2.924% | 5.848% | 8.772% | 73.5923 | 81.3389
P, wt.% | 0.044165 | 0.0013733 | 0.041419 | 0.046912 | 0.040045 | 0.048285 | 3.109% | 6.219% | 9.328% | 0.041957 | 0.046373
Pb, ppm | 37.3356 | 1.6083 | 34.1189 | 40.5522 | 32.5106 | 42.1605 | 4.308% | 8.615% | 12.92% | 35.4688 | 39.2024
Pr, ppm | 1.9902 | 0.20467 | 1.5809 | 2.3996 | 1.3762 | 2.6042 | 10.28% | 20.57% | 30.85% | 1.8907 | 2.0897
Rb, ppm | 18.0996 | 0.94445 | 16.2107 | 19.9885 | 15.2663 | 20.933 | 5.218% | 10.44% | 15.65% | 17.1946 | 19.0046
Re, ppm | 0.0025963 | 0.00054374 | 0.0015088 | 0.0036838 | 0.00096508 | 0.0042275 | 20.94% | 41.89% | 62.83% | 0.0024665 | 0.0027261
S, wt.% | 0.49638 | 0.015859 | 0.46466 | 0.5281 | 0.4488 | 0.54396 | 3.195% | 6.39% | 9.585% | 0.47156 | 0.5212
Sb, ppm | 1.9844 | 0.13859 | 1.7072 | 2.2615 | 1.5686 | 2.4001 | 6.984% | 13.97% | 20.95% | 1.8852 | 2.0836
Sc, ppm | 35.2829 | 1.3985 | 32.4859 | 38.0798 | 31.0875 | 39.4782 | 3.964% | 7.927% | 11.89% | 33.5187 | 37.047
Sm, ppm | 2.6059 | 0.1712 | 2.2635 | 2.9483 | 2.0923 | 3.1195 | 6.57% | 13.14% | 19.71% | 2.4756 | 2.7362
Sn, ppm | 1.3151 | 0.16935 | 0.97635 | 1.6538 | 0.807 | 1.8231 | 12.88% | 25.76% | 38.63% | 1.2493 | 1.3808
Sr, ppm | 155.6961 | 5.9577 | 143.7806 | 167.6115 | 137.8229 | 173.5692 | 3.827% | 7.653% | 11.48% | 147.9113 | 163.4809
Ta, ppm | 0.25106 | 0.023546 | 0.20397 | 0.29815 | 0.18043 | 0.3217 | 9.378% | 18.76% | 28.14% | 0.23851 | 0.26362
Tb, ppm | 0.53307 | 0.030631 | 0.47181 | 0.59434 | 0.44118 | 0.62497 | 5.746% | 11.49% | 17.24% | 0.50642 | 0.55973
Te, ppm | 0.13309 | 0.031364 | 0.070364 | 0.19582 | 0.039 | 0.22718 | 23.57% | 47.13% | 70.7% | 0.12644 | 0.13975
Th, ppm | 1.3048 | 0.12503 | 1.0547 | 1.5548 | 0.92968 | 1.6799 | 9.583% | 19.17% | 28.75% | 1.2395 | 1.37
Ti, wt.% | 0.59234 | 0.020724 | 0.55089 | 0.63379 | 0.53016 | 0.65451 | 3.499% | 6.998% | 10.5% | 0.56272 | 0.62195
Tl, ppm | 0.30994 | 0.017279 | 0.27539 | 0.3445 | 0.25811 | 0.36178 | 5.575% | 11.15% | 16.72% | 0.29445 | 0.32544
Tm, ppm | 0.31763 | 0.028018 | 0.26159 | 0.37366 | 0.23357 | 0.40168 | 8.821% | 17.64% | 26.46% | 0.30175 | 0.33351
U, ppm | 0.39681 | 0.011655 | 0.3735 | 0.42013 | 0.36185 | 0.43178 | 2.937% | 5.874% | 8.812% | 0.37697 | 0.41666
V, ppm | 250.4634 | 8.6231 | 233.2173 | 267.7095 | 224.5943 | 276.3326 | 3.443% | 6.886% | 10.33% | 237.9403 | 262.9866
W, ppm | 38.2764 | 2.524 | 33.2285 | 43.3244 | 30.7046 | 45.8483 | 6.594% | 13.19% | 19.78% | 36.3626 | 40.1903
Y, ppm | 19.5048 | 1.1293 | 17.2461 | 21.7635 | 16.1168 | 22.8928 | 5.79% | 11.58% | 17.37% | 18.5296 | 20.4801
Yb, ppm | 2.1017 | 0.16501 | 1.7717 | 2.4317 | 1.6066 | 2.5967 | 7.852% | 15.7% | 23.55% | 1.9966 | 2.2068
Zn, ppm | 162.9386 | 7.223 | 148.4926 | 177.3847 | 141.2696 | 184.6077 | 4.433% | 8.866% | 13.3% | 154.7917 | 171.0856
Zr, ppm | 57.3284 | 5.581 | 46.1663 | 68.4905 | 40.5853 | 74.0715 | 9.735% | 19.47% | 29.21% | 54.462 | 60.1948
Aqua Regia Digestion
Ag, ppm | 1.7333 | 0.087935 | 1.5575 | 1.9092 | 1.4695 | 1.9971 | 5.073% | 10.15% | 15.22% | 1.6467 | 1.82
Al, wt.% | 3.044 | 0.23623 | 2.5715 | 3.5165 | 2.3353 | 3.7527 | 7.761% | 15.52% | 23.28% | 2.8918 | 3.1962
As, ppm | 85.5163 | 4.0315 | 77.4534 | 93.5793 | 73.4219 | 97.6108 | 4.714% | 9.429% | 14.14% | 81.2405 | 89.7922
B, ppm | 29.9367 | 6.7226 | 16.4915 | 43.3819 | 9.769 | 50.1045 | 22.46% | 44.91% | 67.37% | 28.4399 | 31.4336
Ba, ppm | 40.7607 | 2.0241 | 36.7126 | 44.8089 | 34.6885 | 46.833 | 4.966% | 9.932% | 14.9% | 38.7227 | 42.7988
Be, ppm | 0.25146 | 0.034861 | 0.18174 | 0.32118 | 0.14688 | 0.35604 | 13.86% | 27.73% | 41.59% | 0.23889 | 0.26403
Bi, ppm | 0.062926 | 0.0080993 | 0.046727 | 0.079124 | 0.038628 | 0.087224 | 12.87% | 25.74% | 38.61% | 0.05978 | 0.066072
Ca, wt.% | 2.3084 | 0.22906 | 1.8503 | 2.7665 | 1.6212 | 2.9956 | 9.923% | 19.85% | 29.77% | 2.193 | 2.4238
Cd, ppm | 0.77943 | 0.054069 | 0.67129 | 0.88756 | 0.61722 | 0.94163 | 6.937% | 13.87% | 20.81% | 0.74046 | 0.8184
Ce, ppm | 11.2415 | 0.48243 | 10.2767 | 12.2064 | 9.7942 | 12.6888 | 4.292% | 8.583% | 12.87% | 10.6795 | 11.8036
Co, ppm | 28.7241 | 1.6027 | 25.5188 | 31.9294 | 23.9161 | 33.5321 | 5.579% | 11.16% | 16.74% | 27.2879 | 30.1603
Cr, ppm | 33.7898 | 2.2805 | 29.2288 | 38.3509 | 26.9482 | 40.6314 | 6.749% | 13.5% | 20.25% | 32.1003 | 35.4793
Cs, ppm | 0.8851 | 0.063122 | 0.75886 | 1.0113 | 0.69574 | 1.0745 | 7.132% | 14.26% | 21.39% | 0.84085 | 0.92936
Cu, ppm | 165.3704 | 5.9151 | 153.5403 | 177.2006 | 147.6252 | 183.1157 | 3.577% | 7.154% | 10.73% | 157.1019 | 173.639
Dy, ppm | 2.1191 | 0.31767 | 1.4838 | 2.7545 | 1.1661 | 3.0721 | 14.99% | 29.98% | 44.97% | 2.0132 | 2.2251
Er, ppm | 1.2188 | 0.18469 | 0.84943 | 1.5882 | 0.66474 | 1.7729 | 15.15% | 30.31% | 45.46% | 1.1579 | 1.2798
Eu, ppm | 0.48373 | 0.080263 | 0.32321 | 0.64426 | 0.24294 | 0.72452 | 16.59% | 33.18% | 49.78% | 0.45954 | 0.50792
Fe, wt.% | 5.5679 | 0.17756 | 5.2128 | 5.923 | 5.0352 | 6.1006 | 3.189% | 6.378% | 9.567% | 5.2895 | 5.8463
Ga, ppm | 10.0383 | 0.88101 | 8.2763 | 11.8003 | 7.3953 | 12.6813 | 8.777% | 17.55% | 26.33% | 9.5364 | 10.5402
Gd, ppm | 1.9099 | 0.20134 | 1.5072 | 2.3126 | 1.3059 | 2.5139 | 10.54% | 21.08% | 31.63% | 1.8144 | 2.0054
Ge, ppm | 0.1019 | 0.027089 | 0.047721 | 0.15608 | 0.020632 | 0.18317 | 26.58% | 53.17% | 79.75% | 0.096805 | 0.107
Hf, ppm | 0.45457 | 0.046299 | 0.36198 | 0.54717 | 0.31568 | 0.59347 | 10.19% | 20.37% | 30.56% | 0.43185 | 0.4773
Ho, ppm | 0.4233 | 0.068316 | 0.28667 | 0.55993 | 0.21835 | 0.62825 | 16.14% | 32.28% | 48.42% | 0.40213 | 0.44446
In, ppm | 0.038595 | 0.0026746 | 0.033245 | 0.043944 | 0.030571 | 0.046618 | 6.93% | 13.86% | 20.79% | 0.036665 | 0.040524
K, wt.% | 0.18316 | 0.010759 | 0.16164 | 0.20468 | 0.15088 | 0.21544 | 5.874% | 11.75% | 17.62% | 0.174 | 0.19232
La, ppm | 5.1149 | 0.24459 | 4.6258 | 5.6041 | 4.3812 | 5.8487 | 4.782% | 9.564% | 14.35% | 4.8592 | 5.3707
Li, ppm | 9.8497 | 0.6045 | 8.6407 | 11.0587 | 8.0362 | 11.6632 | 6.137% | 12.27% | 18.41% | 9.3572 | 10.3422
Lu, ppm | 0.14624 | 0.031081 | 0.084083 | 0.20841 | 0.053002 | 0.23949 | 21.25% | 42.51% | 63.76% | 0.13893 | 0.15356
Mg, wt.% | 1.5763 | 0.076432 | 1.4235 | 1.7292 | 1.347 | 1.8056 | 4.849% | 9.697% | 14.55% | 1.4975 | 1.6552
Mn, wt.% | 0.07382 | 0.0049228 | 0.063974 | 0.083665 | 0.059051 | 0.088588 | 6.669% | 13.34% | 20.01% | 0.070129 | 0.077511
Mo, ppm | 2.1628 | 0.090936 | 1.9809 | 2.3446 | 1.89 | 2.4356 | 4.205% | 8.409% | 12.61% | 2.0546 | 2.2709
Na, wt.% | 0.28305 | 0.013249 | 0.25655 | 0.30955 | 0.2433 | 0.3228 | 4.681% | 9.362% | 14.04% | 0.2689 | 0.2972
Nd, ppm | 6.5174 | 0.32728 | 5.8629 | 7.172 | 5.5356 | 7.4993 | 5.022% | 10.04% | 15.06% | 6.1916 | 6.8433
Ni, ppm | 54.7118 | 2.5436 | 49.6245 | 59.799 | 47.0809 | 62.3426 | 4.649% | 9.298% | 13.95% | 51.9762 | 57.4474
P, wt.% | 0.042972 | 0.0019951 | 0.038981 | 0.046962 | 0.036986 | 0.048957 | 4.643% | 9.286% | 13.93% | 0.040823 | 0.04512
Pb, ppm | 37.4093 | 2.1267 | 33.1559 | 41.6626 | 31.0293 | 43.7892 | 5.685% | 11.37% | 17.05% | 35.5388 | 39.2797
Pr, ppm | 1.4185 | 0.069591 | 1.2793 | 1.5577 | 1.2097 | 1.6273 | 4.906% | 9.812% | 14.72% | 1.3476 | 1.4894
Rb, ppm | 7.9145 | 0.52582 | 6.8629 | 8.9661 | 6.337 | 9.4919 | 6.644% | 13.29% | 19.93% | 7.5188 | 8.3102
Re, ppm | 0.00204 | 0.00040089 | 0.0012382 | 0.0028418 | 0.00083732 | 0.0032427 | 19.65% | 39.3% | 58.95% | 0.001938 | 0.002142
S, wt.% | 0.49617 | 0.014666 | 0.46683 | 0.5255 | 0.45217 | 0.54016 | 2.956% | 5.912% | 8.867% | 0.47136 | 0.52097
Sb, ppm | 1.2173 | 0.33293 | 0.55148 | 1.8832 | 0.21854 | 2.2161 | 27.35% | 54.7% | 82.05% | 1.1565 | 1.2782
Sc, ppm | 5.3109 | 0.53039 | 4.2501 | 6.3717 | 3.7198 | 6.9021 | 9.987% | 19.97% | 29.96% | 5.0454 | 5.5765
Sm, ppm | 1.5959 | 0.1825 | 1.2309 | 1.9609 | 1.0484 | 2.1434 | 11.44% | 22.87% | 34.31% | 1.5161 | 1.6757
Sn, ppm | 0.73396 | 0.066904 | 0.60015 | 0.86776 | 0.53324 | 0.93467 | 9.116% | 18.23% | 27.35% | 0.69726 | 0.77065
Sr, ppm | 37.4521 | 4.7073 | 28.0374 | 46.8667 | 23.3301 | 51.574 | 12.57% | 25.14% | 37.71% | 35.5795 | 39.3247
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.31475 | 0.012236 | 0.29028 | 0.33922 | 0.27804 | 0.35146 | 3.888% | 7.775% | 11.66% | 0.29901 | 0.33049
Te, ppm | 0.12598 | 0.01769 | 0.090602 | 0.16136 | 0.072912 | 0.17905 | 14.04% | 28.08% | 42.13% | 0.11968 | 0.13228
Th, ppm | 1.0542 | 0.076633 | 0.90097 | 1.2075 | 0.82433 | 1.2841 | 7.269% | 14.54% | 21.81% | 1.0015 | 1.1069
Ti, wt.% | 0.30233 | 0.047647 | 0.20704 | 0.39762 | 0.15939 | 0.44527 | 15.76% | 31.52% | 47.28% | 0.28721 | 0.31745
Tl, ppm | 0.16316 | 0.0084861 | 0.14618 | 0.18013 | 0.1377 | 0.18861 | 5.201% | 10.4% | 15.6% | 0.155 | 0.17131
Tm, ppm | 0.16728 | 0.020146 | 0.12699 | 0.20757 | 0.10684 | 0.22772 | 12.04% | 24.09% | 36.13% | 0.15891 | 0.17564
U, ppm | 0.25982 | 0.020403 | 0.21901 | 0.30062 | 0.19861 | 0.32103 | 7.853% | 15.71% | 23.56% | 0.24683 | 0.27281
V, ppm | 131.1176 | 14.2955 | 102.5266 | 159.7086 | 88.2311 | 174.0041 | 10.9% | 21.81% | 32.71% | 124.5617 | 137.6735
W, ppm | 27.1493 | 4.653 | 17.8432 | 36.4554 | 13.1902 | 41.1084 | 17.14% | 34.28% | 51.42% | 25.7918 | 28.5067
Y, ppm | 10.9312 | 0.70822 | 9.5148 | 12.3476 | 8.8065 | 13.0559 | 6.479% | 12.96% | 19.44% | 10.3846 | 11.4778
Yb, ppm | 1.0321 | 0.13969 | 0.75275 | 1.3115 | 0.61306 | 1.4512 | 13.53% | 27.07% | 40.6% | 0.98052 | 1.0837
Zn, ppm | 149.3803 | 7.4902 | 134.3999 | 164.3607 | 126.9097 | 171.8509 | 5.014% | 10.03% | 15.04% | 141.9113 | 156.8493
Zr, ppm | 15.9799 | 2.3284 | 11.323 | 20.6368 | 8.9946 | 22.9653 | 14.57% | 29.14% | 43.71% | 15.1809 | 16.7789
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 241b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
CNL*AAS | cyanide leach with atomic absorption spectroscopy
CNL*MS | cyanide leach with inductively coupled plasma mass spectroscopy
CNL*OES | cyanide leach with inductively coupled plasma optical emission spectroscopy
CNL*OES/AAS | cyanide leach with AAS or ICP-OES finish
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Text Values:
N.A. | Not Applicable (Lab 1.27)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 241b
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
AGAT Laboratories, Mississauga, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
CRS Laboratories Oy, Kempele, Northern Ostrobothnia, Finland
Gekko Assay Labs, Ballarat, VIC, Australia
Inspectorate (BV), Lima, Peru
Inspectorate Griffith India, Gandhidham, Gujarat, India
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSALABS, Vancouver, BC, Canada
Nagrom, Perth, WA, Australia
Ontario Geological Survey, Sudbury, Ontario, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
Saskatchewan Research Council, Saskatoon, Saskatchewan, Canada
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 241b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response | Upscaled Value
2 | 1 | ANSLu | 1 | 4 | 4 | 238597 | 0.08438 | 7.1606 | 7.3309
2 | 1 | ANSLu | 1 | 16 | 16 | 238598 | 0.084445 | 7.2891 | 7.3377
2 | 1 | ANSLu | 1 | 7 | 7 | 238599 | 0.084102 | 7.0856 | 7.3269
2 | 1 | ANSLu | 1 | 2 | 2 | 238600 | 0.086846 | 7.3931 | 7.3433
2 | 1 | ANSLu | 1 | 20 | 20 | 238601 | 0.085955 | 7.5427 | 7.3513
2 | 1 | ANSLu | 1 | 13 | 13 | 238602 | 0.08647 | 7.2874 | 7.3376
2 | 1 | ANSLu | 1 | 3 | 3 | 238603 | 0.087029 | 7.4192 | 7.3447
2 | 1 | ANSLu | 1 | 6 | 6 | 238604 | 0.083161 | 7.545 | 7.3514
2 | 1 | ANSLu | 1 | 15 | 15 | 238605 | 0.082582 | 7.284 | 7.3375
2 | 1 | ANSLu | 1 | 14 | 14 | 238606 | 0.087138 | 7.5862 | 7.3536
2 | 1 | ANSLu | 1 | 12 | 12 | 238607 | 0.087702 | 7.1618 | 7.3309
2 | 1 | ANSLu | 1 | 8 | 8 | 238608 | 0.085797 | 7.2702 | 7.3367
2 | 1 | ANSLu | 1 | 17 | 17 | 238609 | 0.088388 | 7.4016 | 7.3437
2 | 1 | ANSLu | 1 | 18 | 18 | 238610 | 0.087826 | 7.4466 | 7.3462
2 | 1 | ANSLu | 1 | 9 | 9 | 238611 | 0.086597 | 7.299 | 7.3383
2 | 1 | ANSLu | 1 | 5 | 5 | 238612 | 0.085297 | 7.2906 | 7.3378
2 | 1 | ANSLu | 1 | 11 | 11 | 238613 | 0.083063 | 7.4213 | 7.3448
2 | 1 | ANSLu | 1 | 10 | 10 | 238614 | 0.086467 | 7.1671 | 7.3312
2 | 1 | ANSLu | 1 | 19 | 19 | 238615 | 0.08435 | 7.2606 | 7.3362
2 | 1 | ANSLu | 1 | 1 | 1 | 238616 | 0.083234 | 7.498 | 7.3489
Mean | [6 empty] | 0.085541 | 7.3405 | 7.3405
Median | [7 empty] | 7.2948 | 7.338
SD | [7 empty] | 0.14177 | 0.0075703
RSD | [7 empty] | 1.931% | 0.1031%
RSD Upscale Factor | [7 empty] | 0.053398
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 241b (Certified Value 7.13 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | INAA | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*MS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 30g | 25g | 50g | 50g | 30g | 50g | 30g | 30g | 50g | 50g | 25g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 15g | 30g | 30g | 25g | 30g | 50g | 30g | 30g | 30g | 3
| 1 | 1 | 7.498 | 6.711 | 7.269 | 6.762 | 6.8539 | 7.3453 | 6.92 | 7.04 | 7.42 | 6.88 | 7.16 | 7.116 | 7.2 | > 5 | 7.15 | 7.26 | 7.46 | 7.393 | 7.11 | 7.322 | 6.97 | 6.779 | 6.99 | 7.053 | 6.4649 | 7.164 | 6.9252 | 7.247 | 6.89 | 6.946 | 7.128 | 6.99 | 1
| 1 | 2 | 7.3931 | 6.031 | 7.073 | 6.922 | 6.9073 | 7.2542 | 6.98 | 7.32 | 7.46 | 6.98 | 7.16 | 6.719 | 7.22 | > 5 | 7.21 | 7.33 | 7.52 | 7.483 | 6.93 | 7.074 | 7.04 | 6.993 | 7.06 | 7.057 | 6.4701 | 7.086 | 6.9687 | 7.24 | 6.98 | 6.986 | 7.256 | 7.1 | #N/A
| 1 | 3 | 7.4192 | 6.132 | 7.243 | 7.068 | 7.0246 | 7.4824 | 6.8 | 7.33 | 7.63 | 7.15 | 7.1 | 7.245 | 7.24 | > 5 | 7.12 | 7.42 | 7.32 | 7.369 | 7.17 | 7.005 | 7.01 | 7.084 | 7.02 | 7.222 | 6.5782 | 7.196 | 7.0745 | 7.27 | 6.92 | 6.926 | 7.18 | 6.86 | 16
| 1 | 4 | 7.1606 | 6.649 | 7.218 | 7.064 | 7.0421 | 7.4986 | 6.98 | 7.24 | 7.43 | 6.98 | 7.14 | 7.092 | 7.3 | > 5 | 7.14 | 7.47 | 7.31 | 7.605 | 6.93 | 7.072 | 6.98 | 7.078 | 7.06 | 7.103 | 6.5184 | 7.218 | 7.2117 | 7.341 | 6.87 | 6.903 | 7.129 | 7.07 | 7.1298
| 1 | 5 | 7.2906 | 6.031 | 7.326 | 6.942 | 6.8646 | 7.5457 | 6.99 | 7.28 | 7.24 | 6.93 | 7.16 | 7.126 | 7.28 | > 5 | 7.27 | 7.34 | 7.4 | 7.276 | 6.98 | 7.094 | 7.01 | 6.764 | 7.03 | 7.177 | 6.5386 | 7.174 | 6.9319 | 7.279 | 6.93 | 6.926 | 6.952 | 7.08 | 7
| 1 | 6 | 7.545 | 5.606 | 7.294 | 7.072 | 6.9527 | 7.4486 | 6.88 | 7.23 | 7.22 | 6.91 | 7.23 | 7.023 | 7.22 | > 5 | 7.16 | 7.16 | 7.19 | 7.04 | 6.91 | 7.165 | 7.02 | 7.188 | 7.03 | 7.374 | 6.5247 | 7.176 | 7.1357 | 7.317 | 6.81 | 6.964 | 6.927 | 7.05
|  | 7 | 7.0856
|  | 8 | 7.2702
|  | 9 | 7.299
|  | 10 | 7.1671
|  | 11 | 7.4213
|  | 12 | 7.1618
|  | 13 | 7.2874
|  | 14 | 7.5862
|  | 15 | 7.284
|  | 16 | 7.2891
|  | 17 | 7.4016
|  | 18 | 7.4466
|  | 19 | 7.2606
|  | 20 | 7.5427
| Mean |  | 7.3405 | 6.1933 | 7.2372 | 6.9717 | 6.9408 | 7.4291 | 6.925 | 7.24 | 7.4 | 6.9717 | 7.1583 | 7.0535 | 7.2433 |  | 7.175 | 7.33 | 7.3667 | 7.361 | 7.005 | 7.122 | 7.005 | 6.981 | 7.0317 | 7.1643 | 6.5158 | 7.169 | 7.0413 | 7.2823 | 6.9 | 6.9418 | 7.0953 | 7.025
| Median |  | 7.2948 | 6.0815 | 7.256 | 7.003 | 6.93 | 7.4655 | 6.95 | 7.26 | 7.425 | 6.955 | 7.16 | 7.104 | 7.23 |  | 7.155 | 7.335 | 7.36 | 7.381 | 6.955 | 7.084 | 7.01 | 7.0355 | 7.03 | 7.14 | 6.5216 | 7.175 | 7.0216 | 7.2745 | 6.905 | 6.936 | 7.1285 | 7.06
| Std Dev. |  | 0.14177 | 0.41882 | 0.088867 | 0.12262 | 0.07989 | 0.10887 | 0.074766 | 0.10602 | 0.15218 | 0.095795 | 0.042151 | 0.17896 | 0.038816 |  | 0.055408 | 0.111 | 0.11827 | 0.19276 | 0.10877 | 0.11057 | 0.025884 | 0.17373 | 0.026394 | 0.12258 | 0.042872 | 0.044984 | 0.11813 | 0.039616 | 0.057966 | 0.029896 | 0.12966 | 0.089163
| Rel.Std.Dev. |  | 1.931% | 6.762% | 1.228% | 1.759% | 1.151% | 1.465% | 1.08% | 1.464% | 2.057% | 1.374% | 0.5888% | 2.537% | 0.5359% |  | 0.7722% | 1.514% | 1.605% | 2.619% | 1.553% | 1.552% | 0.3695% | 2.489% | 0.3754% | 1.711% | 0.658% | 0.6275% | 1.678% | 0.544% | 0.8401% | 0.4307% | 1.827% | 1.269%
| PDM3 |  | 2.955% | -13.13% | 1.506% | -2.218% | -2.65% | 4.199% | -2.872% | 1.546% | 3.79% | -2.218% | 0.4004% | -1.07% | 1.593% |  | 0.6342% | 2.808% | 3.322% | 3.243% | -1.75% | -0.1092% | -1.75% | -2.087% | -1.376% | 0.4845% | -8.611% | 0.55% | -1.241% | 2.14% | -3.223% | -2.636% | -0.4832% | -1.47%
| Z-Score (Absolute) |  | NA | 5.05 | 0.93 | 0.59 | 0.77 | 2.03 | 0.86 | 0.95 | 1.87 | 0.59 | 0.48 | 0.12 | 0.97 | NA | 0.58 | 1.46 | 1.67 | 1.64 | 0.4 | 0.27 | 0.4 | 0.54 | 0.24 | 0.52 | 3.2 | 0.54 | 0.19 | 1.19 | 1 | 0.76 | 0.12 | 0.28
| Analytical results for Pd in OREAS 241b (Indicative Value 11.5 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 1
|  |  | 25g | [31 empty] | 1
| 1 | 1 | 10.8063 | [31 empty] | 1
| 1 | 2 | 11.2277 | [31 empty] | 1
| 1 | 3 | 12.2861 | [31 empty] | 16
| 1 | 4 | 11.4354 | [31 empty] | 11.4675
| 1 | 5 | 11.2842 | [31 empty] | 7
| 1 | 6 | 11.7653
| Mean |  | 11.4675
| Median |  | 11.3598
| Std Dev. |  | 0.50757
| Rel.Std.Dev. |  | 4.426%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pt in OREAS 241b (Indicative Value 9.36 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | Lab | [31 empty] | 1
| No | No | 1.28 | [31 empty] | ppb
|  |  | FA*MS | [31 empty] | 2
|  |  | 25g | [31 empty] | 2
| 1 | 1 | 9.2107 | [31 empty] | 1
| 1 | 2 | 9.3708 | [31 empty] | 1
| 1 | 3 | 9.422 | [31 empty] | 16
| 1 | 4 | 9.3237 | [31 empty] | 9.3569
| 1 | 5 | 9.3317 | [31 empty] | 7
| 1 | 6 | 9.4826
| Mean |  | 9.3569
| Median |  | 9.3513
| Std Dev. |  | 0.093094
| Rel.Std.Dev. |  | 0.9949%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 241b (Certified Value 6.96 ppm) | [25 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.30 | 1.33 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | 2
|  |  | 0.085g | 30g | 25g | 25g | 30g | 50g | 30g | 15g | 50g | 30g | 10g | 30g | 25g | 25g | 25g | 25g | 15g | 25g | 15g | 50g | 25g | 25g | 10g | 20g | 3
| 1 | 1 | 7.498 | 6.95 | 7.73 | 6.82 | 7.0932 | 6.83 | 6.56 | 7.4 | 7.1064 | 6.82 | > 2 | 7.1364 | 7.49 | 7 | 6.56 | 7.01 | 6.7671 | 6.986 | 7.2649 | 6.6654 | 7.255 | 6.5 | 7.083 | 7.1227 | 1
| 1 | 2 | 7.3931 | 6.66 | 7.73 | 6.78 | 6.9907 | 6.97 | 6.81 | 7.5 | 6.5232 | 6.77 | > 2 | 7.0104 | 7.94 | 6.79 | 6.78 | 6.98 | 7.1791 | 7.026 | 7.2231 | 6.6107 | 7.095 | 6.3 | 7.094 | 6.8926 | #N/A
| 1 | 3 | 7.4192 | 6.93 | 7.67 | 6.92 | 6.8335 | 6.76 | 6.65 | 7.31 | 6.5772 | 6.66 | > 2 | 7.1231 | 6.9 | 6.93 | 6.71 | 7 | 6.8392 | 6.932 | 7.4672 | 6.7977 | 7.155 | 6.54 | 7.117 | 6.9788 | 16
| 1 | 4 | 7.1606 | 7.49 | 7.65 | 6.94 | 6.912 | 6.95 | 6.6 | 7.53 | 6.6204 | 6.78 | > 2 | 7.0868 | 7.86 | 6.98 | 6.73 | 6.93 | 7.0658 | 7.035 | 7.1495 | 6.8808 | 7.065 | 6.37 | 7.141 | 6.8091 | 6.9603
| 1 | 5 | 7.2906 | 7.01 | 7.18 | 6.88 | 6.8358 | 6.83 | 6.97 | 7.21 | 6.696 | 6.86 | > 2 | 7.0364 | 7.48 | 6.96 | 6.77 | 6.68 | 6.8598 | 6.947 | 7.7959 | 6.7409 | 7.065 | 6.25 | 7.087 | 7.0469 | 9
| 1 | 6 | 7.545 | 7.32 | 7.64 | 6.87 | 6.7834 | 6.84 | 7.05 | 7.33 | 6.642 | 6.89 | > 2 | 7.0364 | 7.8 | 7.14 | 6.61 | 7 | 6.9319 | 6.992 | 7.4847 | 6.8236 | 7.485 | 6.31 | 7.125 | 6.8696
|  | 7 | 7.0856
|  | 8 | 7.2702
|  | 9 | 7.299
|  | 10 | 7.1671
|  | 11 | 7.4213
|  | 12 | 7.1618
|  | 13 | 7.2874
|  | 14 | 7.5862
|  | 15 | 7.284
|  | 16 | 7.2891
|  | 17 | 7.4016
|  | 18 | 7.4466
|  | 19 | 7.2606
|  | 20 | 7.5427
| Mean |  | 7.3405 | 7.06 | 7.6 | 6.8683 | 6.9081 | 6.8633 | 6.7733 | 7.38 | 6.6942 | 6.7967 |  | 7.0716 | 7.5783 | 6.9667 | 6.6933 | 6.9333 | 6.9405 | 6.9863 | 7.3975 | 6.7532 | 7.1867 | 6.3783 | 7.1078 | 6.9533
| Median |  | 7.2948 | 6.98 | 7.66 | 6.875 | 6.8739 | 6.835 | 6.73 | 7.365 | 6.6312 | 6.8 |  | 7.0616 | 7.645 | 6.97 | 6.72 | 6.99 | 6.8958 | 6.989 | 7.366 | 6.7693 | 7.125 | 6.34 | 7.1055 | 6.9357
| Std Dev. |  | 0.14177 | 0.29799 | 0.20938 | 0.060139 | 0.11614 | 0.080416 | 0.20363 | 0.12133 | 0.21024 | 0.081158 |  | 0.051586 | 0.3839 | 0.11308 | 0.089144 | 0.12738 | 0.15463 | 0.041166 | 0.23693 | 0.10138 | 0.1629 | 0.11686 | 0.023327 | 0.11798
| Rel.Std.Dev. |  | 1.931% | 4.221% | 2.755% | 0.8756% | 1.681% | 1.172% | 3.006% | 1.644% | 3.141% | 1.194% |  | 0.7295% | 5.066% | 1.623% | 1.332% | 1.837% | 2.228% | 0.5892% | 3.203% | 1.501% | 2.267% | 1.832% | 0.3282% | 1.697%
| PDM3 |  | 5.462% | 1.433% | 9.191% | -1.321% | -0.7499% | -1.393% | -2.686% | 6.03% | -3.823% | -2.351% |  | 1.599% | 8.88% | 0.09159% | -3.835% | -0.3873% | -0.2846% | 0.3741% | 6.282% | -2.976% | 3.252% | -8.361% | 2.12% | -0.1007%
| Z-Score (Absolute) |  | NA | 0.49 | 2.83 | 0.34 | 0.17 | 0.36 | 0.75 | 1.88 | 1.1 | 0.65 | NA | 0.54 | 2.74 | 0.09 | 1.1 | 0.06 | 0.03 | 0.17 | 1.95 | 0.84 | 1.04 | 2.46 | 0.7 | 0.03
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## CNL
| Analytical results for Au in OREAS 241b (Certified Value 6.87 ppm) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.30 | 1.32 | ppm
|  |  | INAA | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*OES/AAS | CNL*MS | CNL*OES | CNL*AAS | CNL*AAS | CNL*MS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*AAS | CNL*OES | CNL*AAS | 2
|  |  | 0.085g | 200g | 200g | 200g | 200g | 200g | 50g | 50g | 50g | 05g | 50g | 60g | 30g | 30g | 30g | 200g | 30g | 200g | 200g | 20g | 30g | 05g | 3
| 1 | 1 | 7.498 | 5.656 | 6.9 | 6.65 | 6.724 | 6.6466 | 6.4 | 7.19 | 6.91 | 6.7 | 6.9463 | 6.9 | 7.34 | 6.75 | 7.25 | 6.892 | 6.81 | 6.5 | 6.8645 | 6.81 | 6.785 | 6.93 | 1
| 1 | 2 | 7.3931 | 5.843 | 6.97 | 6.95 | 6.94 | 6.7179 | 6.19 | 7.15 | 6.88 | 6.9 | 6.8955 | 6.92 | 7.32 | 6.87 | 7.16 | 6.819 | 6.92 | 6.53 | 6.8602 | 6.81 | 6.805 | 6.89 | #N/A
| 1 | 3 | 7.4192 | 5.635 | 6.98 | 7.05 | 6.86 | 6.6054 | 6.1 | 7.42 | 7.1 | 6.7 | 6.9035 | 6.91 | 7.47 | 6.83 | 7.21 | 6.843 | 6.86 | 6.56 | 6.8721 | 6.92 | 6.817 | 6.87 | 16
| 1 | 4 | 7.1606 | 5.837 | 6.98 | 7 | 6.872 | 6.6106 | 6.07 | 7.22 | 6.96 | 6.7 | 6.9635 | 6.88 | 7.29 | 6.78 | 7.07 | 6.912 | 6.88 | 6.48 | 6.8427 | 6.72 | 6.903 | 6.9 | 6.8738
| 1 | 5 | 7.2906 | 5.645 | 7.01 | 7.1 | 6.888 | 6.7153 | 6.28 | 7.17 | 6.99 | 6.7 | 6.922 | 6.86 | 7.37 | 6.91 | 7.22 | 6.864 | 6.9 | 6.57 | 6.7917 | 6.96 | 6.727 | 6.89 | 11
| 1 | 6 | 7.545 | 5.788 | 6.95 | 5.275 | 6.904 | 6.8334 | 6.17 | 6.99 | 6.77 | 6.7 | 6.9124 | 6.78 | 7.27 | 6.89 | 7.22 | 6.864 | 6.82 | 6.46 | 6.9033 | 6.75 | 6.75 | 6.88
|  | 7 | 7.0856
|  | 8 | 7.2702
|  | 9 | 7.299
|  | 10 | 7.1671
|  | 11 | 7.4213
|  | 12 | 7.1618
|  | 13 | 7.2874
|  | 14 | 7.5862
|  | 15 | 7.284
|  | 16 | 7.2891
|  | 17 | 7.4016
|  | 18 | 7.4466
|  | 19 | 7.2606
|  | 20 | 7.5427
| Mean |  | 7.3405 | 5.734 | 6.965 | 6.6708 | 6.8647 | 6.6882 | 6.2017 | 7.19 | 6.935 | 6.7333 | 6.9239 | 6.875 | 7.3433 | 6.8383 | 7.1883 | 6.8657 | 6.865 | 6.5167 | 6.8557 | 6.8283 | 6.7978 | 6.8933
| Median |  | 7.2948 | 5.722 | 6.975 | 6.975 | 6.88 | 6.681 | 6.18 | 7.18 | 6.935 | 6.7 | 6.9172 | 6.89 | 7.33 | 6.85 | 7.215 | 6.864 | 6.87 | 6.515 | 6.8624 | 6.81 | 6.795 | 6.89
| Std Dev. |  | 0.14177 | 0.099209 | 0.037283 | 0.70186 | 0.074336 | 0.086389 | 0.12189 | 0.13842 | 0.11113 | 0.08165 | 0.026182 | 0.051284 | 0.071461 | 0.063377 | 0.064936 | 0.033279 | 0.043704 | 0.044121 | 0.037134 | 0.09411 | 0.061561 | 0.020656
| Rel.Std.Dev. |  | 1.931% | 1.73% | 0.5353% | 10.52% | 1.083% | 1.292% | 1.965% | 1.925% | 1.602% | 1.213% | 0.3781% | 0.7459% | 0.9731% | 0.9268% | 0.9034% | 0.4847% | 0.6366% | 0.677% | 0.5416% | 1.378% | 0.9056% | 0.2997%
| PDM3 |  | 6.789% | -16.58% | 1.327% | -2.953% | -0.1328% | -2.7% | -9.778% | 4.6% | 0.8904% | -2.043% | 0.7288% | 0.01751% | 6.831% | -0.5159% | 4.576% | -0.1183% | -0.128% | -5.196% | -0.2626% | -0.6614% | -1.105% | 0.2842%
| Z-Score (Absolute) |  | NA | 10.84 | 0.96 | 1.86 | 0 | 1.69 | 6.36 | 3.12 | 0.67 | 1.26 | 0.57 | 0.1 | 4.59 | 0.25 | 3.1 | 0.01 | 0 | 3.34 | 0.09 | 0.35 | 0.64 | 0.27
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 241b (Certified Value 1.72 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.4 | 1.73 | 1.7385 | 1.8 | 2.15 | 1.7 | 1.71 | 1.8 | 1.6 | 1.75 | 1.95 | 1.81 | 1.73 | 1.8 | 1.6 | 1.5 | 1.79 | 0.28 | 2 | 1.76 | 1.6063 | 1.69 | 1.79 | 1.67 | 1.66 |  |  |  |  | 1
| 1 | 2 | 2.5 |  | 1.7334 | 1.8 | 2.16 | 1.7 | 1.83 | 1.9 | 1.6 | 1.79 | 1.78 | 1.74 | 1.84 | 1.7 | 1.7 | 1.6 | 1.74 | 0.35 | 2 | 1.62 | 1.5875 | 1.7 | 1.73 | 1.73 | 1.63 |  |  |  |  | 18
| 1 | 3 | 2.54 | 1.71 | 1.7073 | 1.6 | 1.97 | 1.7 | 1.79 | 1.9 | 1.7 | 1.75 | 1.78 | 1.76 | 1.74 | 1.7 | 1.7 | 1.6 | 1.7 | 0.3 | 2 | 1.77 | 1.5864 | 1.74 | 1.72 | 1.69 | 1.6 |  |  |  |  | 16
| 1 | 4 | 2.56 | 1.71 | 1.7276 | 1.8 | 1.98 | 1.6 | 1.73 | 1.8 | 1.7 | 1.79 | 1.94 | 1.68 | 1.7 | 1.8 | 1.6 | 1.5 | 1.65 | 0.33 | 2 | 1.69 | 1.5849 | 1.68 | 1.76 | 1.66 | 1.63 |  |  |  |  | 1.7225
| 1 | 5 | 2.59 | 1.77 | 1.7082 | 1.8 | 1.98 | 1.7 | 1.8 | 1.9 | 1.7 | 1.77 | 2.02 | 1.76 | 1.71 | 1.8 | 1.8 | 1.6 | 1.74 | 0.4 | 2 | 1.84 | 1.5508 | 1.7 | 1.81 | 1.67 | 1.68 |  |  |  |  | 13
| 1 | 6 | 2.6 | 1.77 | 1.7225 | 1.7 | 1.85 | 1.6 | 1.73 | 1.9 | 1.7 | 1.75 | 1.96 | 1.8 | 1.63 | 1.8 | 1.7 | 1.6 | 1.7 | 0.5 | 2 | 1.8 | 1.5924 | 1.66 | 1.86 | 1.66 | 1.72
| Mean |  | 2.5317 | 1.738 | 1.7229 | 1.75 | 2.015 | 1.6667 | 1.765 | 1.8667 | 1.6667 | 1.7667 | 1.905 | 1.7583 | 1.725 | 1.7667 | 1.6833 | 1.5667 | 1.72 | 0.36 | 2 | 1.7467 | 1.5847 | 1.695 | 1.7783 | 1.68 | 1.6533
| Median |  | 2.55 | 1.73 | 1.7251 | 1.8 | 1.98 | 1.7 | 1.76 | 1.9 | 1.7 | 1.76 | 1.945 | 1.76 | 1.72 | 1.8 | 1.7 | 1.6 | 1.72 | 0.34 | 2 | 1.765 | 1.5869 | 1.695 | 1.775 | 1.67 | 1.645
| Std Dev. |  | 0.073869 | 0.030332 | 0.012929 | 0.083666 | 0.11912 | 0.05164 | 0.048062 | 0.05164 | 0.05164 | 0.019664 | 0.10075 | 0.046655 | 0.068337 | 0.05164 | 0.075277 | 0.05164 | 0.047749 | 0.08025 | 0 | 0.079415 | 0.018362 | 0.026646 | 0.052694 | 0.026833 | 0.04274
| Rel.Std.Dev. |  | 2.918% | 1.745% | 0.7504% | 4.781% | 5.912% | 3.098% | 2.723% | 2.766% | 3.098% | 1.113% | 5.289% | 2.653% | 3.962% | 2.923% | 4.472% | 3.296% | 2.776% | 22.29% | 0% | 4.547% | 1.159% | 1.572% | 2.963% | 1.597% | 2.585%
| PDM3 |  | 46.98% | 0.8995% | 0.02403% | 1.596% | 16.98% | -3.242% | 2.467% | 8.369% | -3.242% | 2.564% | 10.59% | 2.08% | 0.1448% | 2.564% | -2.274% | -9.047% | -0.1455% | -79.1% | 16.11% | 1.403% | -7.999% | -1.597% | 3.241% | -2.468% | -4.016%
| Z-Score (Absolute) |  | 11.36 | 0.09 | 0.12 | 0.26 | 4.02 | 0.92 | 0.48 | 1.92 | 0.92 | 0.5 | 2.46 | 0.38 | 0.09 | 0.5 | 0.68 | 2.34 | 0.16 | 19.47 | NA | 0.22 | 2.08 | 0.52 | 0.66 | 0.73 | 1.11
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Al in OREAS 241b (Certified Value 6.51 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 6.94 | 6.2724 | 6.6407 | 6.4767 | 6.6 | 7.145 | 6.45 | 6.0941 | 6.15 | 6.61 | 6.48 | 7.11 | 6.34 | 6.37 | 6.49 | 6.606 | 6.42 | 6.47 | 4.28 | 6.545 | 6.517 | 6.42 | 6.06 | 6.41 | 6.97 | 6.7 | 6.4 |  |  | 1
| 1 | 2 | 7.07 | 6.4663 | 6.6904 | 6.7153 | 6.38 | 7.304 | 6.58 | 6.3381 | 6.17 | 6.51 | 6.42 | 6.7 | 6.29 | 6.29 | 6.21 | 6.538 | 6.22 | 6.62 | 3.99 | 6.4008 | 6.456 | 6.26 | 6.096 | 6.33 | 6.87 | 6.67 | 6.6 |  |  | #N/A
| 1 | 3 | 6.99 | 6.3116 | 6.649 | 6.4342 | 6.37 | 7.304 | 6.76 | 6.3042 | 6.21 | 6.57 | 6.7 | 6.67 | 6.24 | 6.38 | 6.37 | 6.603 | 6.26 | 6.38 | 4.32 | 6.5143 | 6.508 | 6.54 | 6.342 | 6.66 | 6.92 | 6.83 | 6.4 |  |  | 16
| 1 | 4 | 6.56 | 6.2859 | 6.6682 | 6.4702 | 6.46 | 6.88 | 6.57 | 6.1508 | 6.14 | 6.58 | 6.73 | 6.92 | 6.19 | 6.28 | 6.64 | 6.453 | 6.5 | 6.51 | 4.35 | 6.4641 | 6.434 | 6.46 | 6.23 | 6.52 | 6.78 | 6.72 | 6.6 |  |  | 6.5128
| 1 | 5 | 7.04 | 6.3657 | 6.6679 | 6.4858 | 6.55 | 7.092 | 6.64 | 6.3561 | 6.12 | 6.62 | 6.51 | 7.05 | 6.92 | 6.34 | 6.53 | 6.496 | 6.49 | 6.37 | 3.84 | 6.3543 | 6.438 | 6.59 | 6.145 | 6.59 | 7.02 | 6.69 | 7.2 |  |  | 14
| 1 | 6 | 6.93 | 6.27 | 6.6529 | 6.6146 | 6.46 | 7.145 | 6.77 | 6.3362 | 6.13 | 6.67 | 6.49 | 6.86 | 6.26 | 6.44 | 6.61 | 6.434 | 6.2 | 6.67 | 4.13 | 6.3834 | 6.437 | 6.43 | 6.098 | 6.65 | 6.93 | 6.75 | 7.3
| Mean |  | 6.9217 | 6.3286 | 6.6615 | 6.5328 | 6.47 | 7.145 | 6.6283 | 6.2633 | 6.1533 | 6.5933 | 6.555 | 6.885 | 6.3733 | 6.35 | 6.475 | 6.5217 | 6.3483 | 6.5033 | 4.1517 | 6.4437 | 6.465 | 6.45 | 6.1618 | 6.5267 | 6.915 | 6.7267 | 6.75
| Median |  | 6.965 | 6.2988 | 6.6604 | 6.4812 | 6.46 | 7.145 | 6.61 | 6.3202 | 6.145 | 6.595 | 6.5 | 6.89 | 6.275 | 6.355 | 6.51 | 6.517 | 6.34 | 6.49 | 4.205 | 6.4324 | 6.447 | 6.445 | 6.1215 | 6.555 | 6.925 | 6.71 | 6.6
| Std Dev. |  | 0.18541 | 0.076213 | 0.017804 | 0.10861 | 0.091214 | 0.15722 | 0.12254 | 0.11178 | 0.03266 | 0.053914 | 0.12787 | 0.17897 | 0.27245 | 0.06 | 0.16121 | 0.073557 | 0.13747 | 0.12291 | 0.2039 | 0.07633 | 0.037704 | 0.11419 | 0.10599 | 0.13367 | 0.083126 | 0.057504 | 0.39875
| Rel.Std.Dev. |  | 2.679% | 1.204% | 0.2673% | 1.663% | 1.41% | 2.2% | 1.849% | 1.785% | 0.5308% | 0.8177% | 1.951% | 2.599% | 4.275% | 0.9449% | 2.49% | 1.128% | 2.165% | 1.89% | 4.911% | 1.185% | 0.5832% | 1.77% | 1.72% | 2.048% | 1.202% | 0.8549% | 5.907%
| PDM3 |  | 6.278% | -2.827% | 2.283% | 0.307% | -0.6571% | 9.707% | 1.774% | -3.832% | -5.519% | 1.237% | 0.648% | 5.715% | -2.141% | -2.5% | -0.5803% | 0.1362% | -2.525% | -0.1453% | -36.25% | -1.062% | -0.7339% | -0.9642% | -5.389% | 0.213% | 6.176% | 3.284% | 3.642%
| Z-Score (Absolute) |  | 1.84 | 0.77 | 0.7 | 0.13 | 0.15 | 2.82 | 0.55 | 1.06 | 1.54 | 0.4 | 0.23 | 1.68 | 0.57 | 0.67 | 0.12 | 0.08 | 0.68 | 0 | 10.34 | 0.26 | 0.17 | 0.23 | 1.5 | 0.1 | 1.81 | 0.98 | 1.08
| Analytical results for As in OREAS 241b (Certified Value 87 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 0
1
| 1 | 1 | 96 | 87.1 | 90.4987 | 65.4 | 86 | 83.4 | 84 | 84.9 | 86.8 | 93.3 | 87.8 | 84.6 | 88 | 89 | 87 | 84 | 56.59 | 84 | 86.5 | 70.8276 | 92.1 | 84 | 77.6 | 84.3 | [5 empty] | 1
| 1 | 2 | 100 | 87 | 93.9071 | 69.3 | 86 | 87.9 | 84 | 86.4 | 89.4 | 88.9 | 87 | 88.9 | 83 | 88 | 80 | 89 | 64.96 | 83 | 82.9 | 71.1671 | 92 | 86 | 77.4 | 81.6 | [5 empty] | 19
| 1 | 3 | 100 | 87.3 | 90.5868 | 54.3 | 86 | 86.4 | 83 | 86.7 | 86.6 | 86.2 | 86.3 | 86.9 | 84 | 92 | 81 | 87 | 53.8 | 82 | 88 | 70.9157 | 90.2 | 87 | 77.4 | 79.1 | [5 empty] | 16
| 1 | 4 | 98 | 87.2 | 91.0304 | 56.8 | 86 | 84.4 | 82 | 86.1 | 86.9 | 91 | 86 | 84.9 | 90 | 93 | 85 | 86 | 55.89 | 86 | 84.5 | 74.6707 | 90.2 | 86 | 74.1 | 77.9 | [5 empty] | 86.8202
| 1 | 5 | 102 | 88.2 | 90.5987 | 68 | 89 | 86 | 82 | 83.1 | 85.9 | 95.5 | 91 | 85.9 | 89 | 90 | 85 | 84 | 67.34 | 85 | 88 | 69.1983 | 91.1 | 91 | 76.6 | 82.8 | [5 empty] | 15
| 1 | 6 | 100 | 86.3 | 93.086 | 63.1 | 87 | 85.8 | 82 | 87.2 | 90 | 93.4 | 86.6 | 86.2 | 89 | 89 | 81 | 90 | 94.38 | 86 | 85.2 | 68.3721 | 90.5 | 87 | 73.9 | 90.4
| Mean |  | 99.3333 | 87.1833 | 91.6179 | 62.8167 | 86.6667 | 85.65 | 82.8333 | 85.7333 | 87.6 | 91.3833 | 87.45 | 86.2333 | 87.1667 | 90.1667 | 83.1667 | 86.6667 | 65.4933 | 84.3333 | 85.85 | 70.8586 | 91.0167 | 86.8333 | 76.1667 | 82.6833
| Median |  | 100 | 87.15 | 90.8145 | 64.25 | 86 | 85.9 | 82.5 | 86.25 | 86.85 | 92.15 | 86.8 | 86.05 | 88.5 | 89.5 | 83 | 86.5 | 60.775 | 84.5 | 85.85 | 70.8717 | 90.8 | 86.5 | 77 | 82.2
| Std Dev. |  | 2.0656 | 0.61128 | 1.4897 | 6.0727 | 1.2111 | 1.5745 | 0.98319 | 1.5029 | 1.6745 | 3.402 | 1.848 | 1.5565 | 2.9269 | 1.9408 | 2.8577 | 2.5033 | 15.1417 | 1.633 | 2.0305 | 2.1712 | 0.86583 | 2.3166 | 1.7143 | 4.4513
| Rel.Std.Dev. |  | 2.079% | 0.7011% | 1.626% | 9.667% | 1.397% | 1.838% | 1.187% | 1.753% | 1.912% | 3.723% | 2.113% | 1.805% | 3.358% | 2.152% | 3.436% | 2.888% | 23.12% | 1.936% | 2.365% | 3.064% | 0.9513% | 2.668% | 2.251% | 5.383%
| PDM3 |  | 14.41% | 0.4182% | 5.526% | -27.65% | -0.1769% | -1.348% | -4.592% | -1.252% | 0.8981% | 5.256% | 0.7254% | -0.676% | 0.399% | 3.854% | -4.208% | -0.1769% | -24.56% | -2.864% | -1.118% | -18.38% | 4.833% | 0.01508% | -12.27% | -4.765%
| Z-Score (Absolute) |  | 3.22 | 0.18 | 1.29 | 5.9 | 0.05 | 0.2 | 0.9 | 0.18 | 0.29 | 1.23 | 0.25 | 0.05 | 0.18 | 0.93 | 0.82 | 0.05 | 5.23 | 0.53 | 0.15 | 3.89 | 1.14 | 0.1 | 2.57 | 0.94
| Analytical results for B in OREAS 241b (Indicative Value 35.7 ppm) | [30 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.03 | 1.25 | 1.27 | [26 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 1
1
| 1 | 1 | 53.65 | 17961.22 | 10 | [26 empty] | 1
| 1 | 2 | 53.24 | 23765.86 | 10 | [26 empty] | 3
| 1 | 3 | 144.65 | 18995.97 | 15 | [26 empty] | 16
| 1 | 4 | 71.46 | 19797.88 | 10 | [26 empty] | 35.6707
| 1 | 5 | 56.8 | 30151.89 | 10 | [26 empty] | 9
| 1 | 6 | 67.39 | 46017.27 | 10
| Mean |  | 74.5317 | 26115.015 | 10.8333
| Median |  | 62.095 | 21781.87 | 10
| Std Dev. |  | 35.1591 | 10723.6749 | 2.0412
| Rel.Std.Dev. |  | 47.17% | 41.06% | 18.84%
| PDM3 |  | 108.9% | 73111% | -69.63%
| Z-Score (Absolute) |  | 0 | 275.66 | 0.67
| Analytical results for Ba in OREAS 241b (Certified Value 257 ppm) | [30 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 259 | 255.5 | 259.1995 | 255.3973 | 237 | 291 | 250 | 248 | 241 | 261 | 260 | 280 | 250 | 260 | 205 | 256 | 230 | 240 | 695.96 | 250.083 | 260 | 253.435 | 260 | 256.5803 | 254 | 282 | 247 | 279 |  | 1
| 1 | 2 | 260 | 261.9 | 261.2008 | 266.4631 | 233 | 291 | 250 | 261.7 | 236 | 261 | 260 | 270 | 250 | 260 | 193 | 252 | 229 | 244 | 1253.09 | 248.859 | 255 | 256.44 | 250 | 256.5901 | 251 | 281 | 245 | 271 |  | 20
| 1 | 3 | 261 | 261.5 | 262.856 | 253.2443 | 232 | 298 | 253 | 278.3 | 244 | 259 | 270 | 270 | 250 | 260 | 199 | 255 | 231 | 236 | 1514.11 | 248.463 | 255 | 259.2778 | 260 | 257.9613 | 263 | 287 | 250 | 265 |  | 16
| 1 | 4 | 259 | 252.8 | 264.8439 | 256.9392 | 233 | 281 | 252 | 253 | 240 | 259 | 270 | 280 | 250 | 250 | 209 | 250 | 232 | 240 | 1434.25 | 246.24 | 255 | 252.7459 | 260 | 255.9853 | 259 | 281 | 249 | 271 |  | 256.9674
| 1 | 5 | 260 | 261.5 | 272.6822 | 255.9307 | 230 | 289 | 259 | 263.8 | 236 | 259 | 260 | 280 | 280 | 260 | 207 | 266 | 234 | 236 | 1881.95 | 244.062 | 255 | 254.4144 | 260 | 248.5321 | 261 | 296 | 248 | 284 |  | 16
| 1 | 6 | 258 | 253.5 | 267.7897 | 261.7166 | 244 | 292 | 258 | 260.3 | 243 | 263 | 260 | 270 | 250 | 260 | 209 | 264 | 232 | 246 | 1345.56 | 244.206 | 255 | 253.5197 | 260 | 256.9557 | 265 | 302 | 250 | 301
| Mean |  | 259.5 | 257.7833 | 264.762 | 258.2819 | 234.8333 | 290.3333 | 253.6667 | 260.85 | 240 | 260.3333 | 263.3333 | 275 | 255 | 258.3333 | 203.6667 | 257.1667 | 231.3333 | 240.3333 | 1354.1533 | 246.9855 | 255.8333 | 254.9721 | 258.3333 | 255.4341 | 258.8333 | 288.1667 | 248.1667 | 278.5
| Median |  | 259.5 | 258.5 | 263.85 | 256.435 | 233 | 291 | 252.5 | 261 | 240.5 | 260 | 260 | 275 | 250 | 260 | 206 | 255.5 | 231.5 | 240 | 1389.905 | 247.3515 | 255 | 253.967 | 260 | 256.5852 | 260 | 284.5 | 248.5 | 275
| Std Dev. |  | 1.0488 | 4.312 | 4.8818 | 4.894 | 5.0365 | 5.5015 | 3.9328 | 10.4017 | 3.4059 | 1.633 | 5.164 | 5.4772 | 12.2474 | 4.0825 | 6.4083 | 6.4627 | 1.7512 | 4.0825 | 388.3496 | 2.5344 | 2.0412 | 2.4671 | 4.0825 | 3.4436 | 5.3821 | 8.8863 | 1.9408 | 12.8957
| Rel.Std.Dev. |  | 0.4042% | 1.673% | 1.844% | 1.895% | 2.145% | 1.895% | 1.55% | 3.988% | 1.419% | 0.6273% | 1.961% | 1.992% | 4.803% | 1.58% | 3.146% | 2.513% | 0.757% | 1.699% | 28.68% | 1.026% | 0.7979% | 0.9676% | 1.58% | 1.348% | 2.079% | 3.084% | 0.7821% | 4.63%
| PDM3 |  | 0.9856% | 0.3175% | 3.033% | 0.5115% | -8.614% | 12.98% | -1.284% | 1.511% | -6.603% | 1.31% | 2.477% | 7.017% | -0.7656% | 0.5316% | -20.74% | 0.07755% | -9.976% | -6.473% | 427% | -3.884% | -0.4413% | -0.7765% | 0.5316% | -0.5967% | 0.7261% | 12.14% | -3.425% | 8.38%
| Z-Score (Absolute) |  | 0.2 | 0.03 | 0.94 | 0.03 | 3.24 | 4.51 | 0.61 | 0.39 | 2.52 | 0.32 | 0.74 | 2.37 | 0.42 | 0.04 | 7.58 | 0.12 | 3.72 | 2.47 | 152.92 | 1.54 | 0.31 | 0.43 | 0.04 | 0.36 | 0.11 | 4.2 | 1.38 | 2.86
| Analytical results for Be in OREAS 241b (Certified Value 0.46 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 0.4 | 0.54 | 0.34 | 0.46683 | < 0.5 | 0.5 | 0.7 | 0.6 | 0.46 | < 5 | < 1 | 0.48 | 0.55 | 0.52 | 0.46 | < 0.5 | 0.5 | < 1 | 0.46 | 0.5 | 0.44693 | 0.45 | 0.38121 | 0.1 | 0.6 | 0.45 | 0.4 |  |  | 1
| 1 | 2 | 0.4 | 0.52 | 0.33 | 0.45485 | 0.50342 | 0.5 | 0.7 | 0.8 | 0.52 | < 5 | < 1 | 0.48 | 0.53 | 0.52 | 0.45 | < 0.5 | 0.5 | < 1 | 0.45 | 0.5 | 0.42647 | 0.44 | 0.38261 | 0.13 | 0.5 | 0.45 | 0.4 |  |  | 21
| 1 | 3 | 0.5 | 0.54 | 0.41 | 0.46444 | < 0.5 | 0.5 | 0.6 | 0.6 | 0.49 | < 5 | < 1 | 0.5 | 0.52 | 0.51 | 0.47 | < 0.5 | 0.5 | < 1 | 0.45 | 0.5 | 0.45667 | 0.47 | 0.39065 | 0.1 | 0.5 | 0.47 | 0.42 |  |  | 16
| 1 | 4 | 0.5 | 0.5 | 0.38 | 0.44766 | < 0.5 | 0.5 | 0.6 | 0.7 | 0.47 | < 5 | < 1 | 0.51 | 0.54 | 0.49 | 0.47 | < 0.5 | 0.5 | < 1 | 0.42 | 0.5 | 0.44062 | 0.46 | 0.37659 | 0.15 | 0.6 | 0.44 | 0.41 |  |  | 0.4644
| 1 | 5 | 0.5 | 0.56 | 0.4 | 0.45129 | < 0.5 | 0.5 | 0.6 | 0.7 | 0.55 | < 5 | < 1 | 0.47 | 0.55 | 0.56 | 0.46 | < 0.5 | 0.5 | < 1 | 0.48 | < 0.5 | 0.46134 | 0.46 | 0.39295 | 0.14 | 0.6 | 0.43 | 0.43 |  |  | 17
| 1 | 6 | 0.5 | 0.6 | 0.39 | 0.44466 | < 0.5 | 0.5 | 0.5 | 0.7 | 0.45 | < 5 | < 1 | 0.48 | 0.53 | 0.5 | 0.47 | < 0.5 | 0.5 | < 1 | 0.47 | < 0.5 | 0.44011 | 0.45 | 0.37918 | 0.11 | 0.7 | 0.42 | 0.42
| Mean |  | 0.46667 | 0.54333 | 0.375 | 0.45495 | 0.50342 | 0.5 | 0.61667 | 0.68333 | 0.49 |  |  | 0.48667 | 0.53667 | 0.51667 | 0.46333 |  | 0.5 |  | 0.455 | 0.5 | 0.44536 | 0.455 | 0.38387 | 0.12167 | 0.58333 | 0.44333 | 0.41333
| Median |  | 0.5 | 0.54 | 0.385 | 0.45307 | 0.50342 | 0.5 | 0.6 | 0.7 | 0.48 |  |  | 0.48 | 0.535 | 0.515 | 0.465 |  | 0.5 |  | 0.455 | 0.5 | 0.44378 | 0.455 | 0.38191 | 0.12 | 0.6 | 0.445 | 0.415
| Std Dev. |  | 0.05164 | 0.034448 | 0.032711 | 0.008985 |  | 0 | 0.075277 | 0.075277 | 0.038471 |  |  | 0.015055 | 0.012111 | 0.024221 | 0.008165 |  | 0 |  | 0.020736 | 0 | 0.012591 | 0.010488 | 0.0065121 | 0.02137 | 0.075277 | 0.017512 | 0.012111
| Rel.Std.Dev. |  | 11.07% | 6.34% | 8.723% | 1.975% |  | 0% | 12.21% | 11.02% | 7.851% |  |  | 3.094% | 2.257% | 4.688% | 1.762% |  | 0% |  | 4.557% | 0% | 2.827% | 2.305% | 1.696% | 17.56% | 12.9% | 3.95% | 2.93%
| PDM3 |  | 0.4891% | 17% | -19.25% | -2.033% | 8.402% | 7.667% | 32.79% | 47.14% | 5.514% |  |  | 4.796% | 15.56% | 11.26% | -0.2286% |  | 7.667% |  | -2.023% | 7.667% | -4.1% | -2.023% | -17.34% | -73.8% | 25.61% | -4.535% | -11%
| Z-Score (Absolute) |  | NA | 1.32 | 1.2 | 0 | 2.44 | NA | NA | NA | 0.52 | 30.64 | 0.67 | 0.47 | 1.22 | 0.92 | 0.12 | 3.07 | NA | 0.67 | 0 | NA | 0.14 | 0 | 1.07 | 4.99 | NA | 0.17 | 0.62
Results from laboratories 1.01, 1.07, 1.08, 1.10, 1.22, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 241b (Certified Value 0.069 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.13 | 0.07 | 0.089485 | 0.07 | 0.3 | 0.07 | 0.08 | < 0.1 | 0.06 | 0.06 | 0.07 | 0.06 | < 5 | 0.09 | < 0.1 | 0.07 | < 0.1 | 0.1 | < 0.11 | 0.06 | 0.0335 | 0.07 | 0.08 | 0.06 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.19 | 0.07 | 0.09301 | 0.07 | 0.3 | 0.07 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.06 | 0.06 | < 5 | 0.08 | < 0.1 | 0.07 | < 0.1 | < 0.1 | < 0.11 | 0.06 | 0.0306 | 0.06 | 0.08 | 0.06 | 0.05 |  |  |  |  | 22
| 1 | 3 | 0.16 | 0.06 | 0.085975 | 0.06 | 0.2 | 0.08 | 0.07 | < 0.1 | 0.06 | 0.07 | 0.07 | 0.05 | < 5 | 0.08 | < 0.1 | 0.07 | 3.32 | < 0.1 | < 0.11 | 0.06 | 0.0343 | 0.07 | 0.08 | 0.06 | 0.05 |  |  |  |  | 16
| 1 | 4 | 0.07 | 0.07 | 0.084213 | 0.06 | 0.3 | 0.09 | 0.07 | < 0.1 | 0.07 | 0.07 | 0.06 | 0.06 | < 5 | 0.07 | < 0.1 | 0.08 | < 0.1 | 0.2 | < 0.11 | 0.06 | 0.0356 | 0.07 | 0.07 | 0.06 | 0.05 |  |  |  |  | 0.069047
| 1 | 5 | 0.16 | 0.06 | 0.086721 | 0.07 | 0.3 | 0.08 | 0.08 | < 0.1 | 0.07 | 0.06 | 0.07 | 0.06 | < 5 | 0.09 | < 0.1 | 0.07 | < 0.1 | 0.1 | < 0.11 | 0.06 | 0.032 | 0.07 | 0.07 | 0.06 | 0.07 |  |  |  |  | 18
| 1 | 6 | 0.14 | 0.06 | 0.089122 | 0.07 | 0.2 | 0.09 | 0.07 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.06 | < 5 | 0.07 | < 0.1 | 0.08 | < 0.1 | 0.1 | < 0.11 | 0.07 | 0.0314 | 0.07 | 0.09 | 0.05 | 0.06
| Mean |  | 0.14167 | 0.065 | 0.088088 | 0.066667 | 0.26667 | 0.08 | 0.073333 |  | 0.065 | 0.066667 | 0.066667 | 0.058333 |  | 0.08 |  | 0.073333 | 3.32 | 0.125 |  | 0.061667 | 0.0329 | 0.068333 | 0.078333 | 0.058333 | 0.055
| Median |  | 0.15 | 0.065 | 0.087921 | 0.07 | 0.3 | 0.08 | 0.07 |  | 0.065 | 0.07 | 0.07 | 0.06 |  | 0.08 |  | 0.07 | 3.32 | 0.1 |  | 0.06 | 0.03275 | 0.07 | 0.08 | 0.06 | 0.05
| Std Dev. |  | 0.040702 | 0.0054772 | 0.0031176 | 0.005164 | 0.05164 | 0.0089443 | 0.005164 |  | 0.0054772 | 0.005164 | 0.005164 | 0.0040825 |  | 0.0089443 |  | 0.005164 |  | 0.05 |  | 0.0040825 | 0.0018953 | 0.0040825 | 0.0075277 | 0.0040825 | 0.0083666
| Rel.Std.Dev. |  | 28.73% | 8.427% | 3.539% | 7.746% | 19.36% | 11.18% | 7.042% |  | 8.427% | 7.746% | 7.746% | 6.999% |  | 11.18% |  | 7.042% |  | 40% |  | 6.62% | 5.761% | 5.974% | 9.61% | 6.999% | 15.21%
| PDM3 |  | 105.2% | -5.861% | 27.58% | -3.448% | 286.2% | 15.86% | 6.208% |  | -5.861% | -3.448% | -3.448% | -15.52% |  | 15.86% |  | 6.208% | 4708% | 81.04% |  | -10.69% | -52.35% | -1.034% | 13.45% | -15.52% | -20.34%
| Z-Score (Absolute) |  | 4.33 | 0.1 | 1.24 | 0 | NA | 0.77 | 0.39 | 0.96 | 0.1 | 0 | 0 | 0.48 | 140.64 | 0.77 | 0.96 | 0.39 | 30.54 | NA | 0.67 | 0.29 | 1.95 | 0.1 | 0.67 | 0.48 | 0.67
Results from laboratories 1.08 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 241b (Certified Value 5.79 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 5.49 | 5.9374 | 5.7442 | 5.8655 | 5.31 | 6.611 | 5.74 | 5.7154 | 5.48 | 5.79 | 5.64 | 6.56 | 5.8 | 5.8 | 5.71 | 6.027 | 5.56 | 5.861 | 4.43 | 5.7065 | 6.004 | 5.95 | 6.1607 | 5.54 | 5.32 | 5.58 | 6.37 |  |  | 1
| 1 | 2 | 5.58 | 5.982 | 5.7395 | 6.0856 | 5.33 | 6.575 | 5.82 | 6.0749 | 5.42 | 5.81 | 5.64 | 6.23 | 5.73 | 5.72 | 5.45 | 5.934 | 5.48 | 5.974 | 3.8 | 5.6675 | 5.954 | 5.74 | 6.1227 | 5.46 | 5.27 | 5.56 | 6.24 |  |  | #N/A
| 1 | 3 | 5.59 | 6.0095 | 5.7115 | 5.792 | 5.26 | 6.682 | 5.95 | 5.8999 | 5.44 | 5.82 | 5.86 | 6.17 | 5.69 | 5.79 | 5.59 | 6.052 | 5.51 | 5.806 | 5 | 5.7218 | 5.995 | 6.06 | 6.3287 | 5.73 | 5.28 | 5.68 | 6.15 |  |  | 16
| 1 | 4 | 5.68 | 5.9542 | 5.7614 | 5.9202 | 5.25 | 6.239 | 5.8 | 5.8001 | 5.49 | 5.73 | 5.93 | 6.32 | 5.64 | 5.76 | 5.85 | 5.968 | 5.58 | 5.881 | 4.73 | 5.7103 | 5.965 | 5.95 | 6.1147 | 5.63 | 5.22 | 5.6 | 6.27 |  |  | 5.7858
| 1 | 5 | 5.62 | 6.0505 | 5.7514 | 5.8312 | 5.17 | 6.439 | 5.87 | 5.9818 | 5.39 | 5.78 | 5.62 | 6.43 | 6.3 | 5.8 | 5.75 | 5.904 | 5.67 | 5.819 | 3.83 | 5.6201 | 5.955 | 6.09 | 6.2767 | 5.69 | 5.36 | 5.59 | 6.17 |  |  | 19
| 1 | 6 | 5.59 | 5.9494 | 5.7325 | 5.9634 | 5.55 | 6.547 | 6.01 | 6.0442 | 5.48 | 5.82 | 5.52 | 6.3 | 5.73 | 5.83 | 5.74 | 5.877 | 5.51 | 5.992 | 2.45 | 5.6496 | 5.962 | 5.8 | 6.0957 | 5.72 | 5.34 | 5.64 | 6.07
| Mean |  | 5.5917 | 5.9805 | 5.7401 | 5.9096 | 5.3117 | 6.5155 | 5.865 | 5.9194 | 5.45 | 5.7917 | 5.7017 | 6.335 | 5.815 | 5.7833 | 5.6817 | 5.9603 | 5.5517 | 5.8888 | 4.04 | 5.6793 | 5.9725 | 5.9317 | 6.1832 | 5.6283 | 5.2983 | 5.6083 | 6.2117
| Median |  | 5.59 | 5.9681 | 5.7419 | 5.8928 | 5.285 | 6.561 | 5.845 | 5.9409 | 5.46 | 5.8 | 5.64 | 6.31 | 5.73 | 5.795 | 5.725 | 5.951 | 5.535 | 5.871 | 4.13 | 5.687 | 5.9635 | 5.95 | 6.1417 | 5.66 | 5.3 | 5.595 | 6.205
| Std Dev. |  | 0.061779 | 0.043034 | 0.017158 | 0.10572 | 0.12937 | 0.15728 | 0.10015 | 0.14133 | 0.04 | 0.034303 | 0.15779 | 0.14096 | 0.24337 | 0.038297 | 0.14091 | 0.068873 | 0.068532 | 0.078065 | 0.91429 | 0.040083 | 0.02151 | 0.13877 | 0.096357 | 0.10834 | 0.051543 | 0.044008 | 0.10477
| Rel.Std.Dev. |  | 1.105% | 0.7196% | 0.2989% | 1.789% | 2.436% | 2.414% | 1.708% | 2.388% | 0.7339% | 0.5923% | 2.767% | 2.225% | 4.185% | 0.6622% | 2.48% | 1.156% | 1.234% | 1.326% | 22.63% | 0.7058% | 0.3602% | 2.339% | 1.558% | 1.925% | 0.9728% | 0.7847% | 1.687%
| PDM3 |  | -3.356% | 3.365% | -0.7903% | 2.14% | -8.195% | 12.61% | 1.368% | 2.308% | -5.804% | 0.1009% | -1.455% | 9.492% | 0.5042% | -0.04316% | -1.8% | 3.016% | -4.047% | 1.78% | -30.17% | -1.841% | 3.226% | 2.521% | 6.868% | -2.722% | -8.426% | -3.068% | 7.36%
| Z-Score (Absolute) |  | 0.8 | 0.75 | 0.21 | 0.47 | 1.92 | 2.89 | 0.29 | 0.51 | 1.37 | 0 | 0.36 | 2.17 | 0.09 | 0.03 | 0.44 | 0.67 | 0.96 | 0.39 | 7 | 0.45 | 0.72 | 0.56 | 1.57 | 0.65 | 1.97 | 0.73 | 1.68
| Analytical results for Cd in OREAS 241b (Certified Value 0.78 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 0.92 | 0.78 | 0.29 | 0.86006 | 0.6 | 1.1 | 0.78 | 0.86 | 1.2 | 0.8 | 0.79 | 0.83 | 0.74 | 0.79 | < 1 | 0.82 | 0.7 | 0.76 | < 0.1 | 1 | 0.74088 | 0.77 | 0.759 | 0.79 | 0.86 | 0.79 | 0.82 |  |  | 1
| 1 | 2 | 0.9 | 0.73 | 0.28 | 0.91047 | 0.6 | 1.1 | 0.78 | 0.76 | 1.2 | 0.9 | 0.78 | 0.8 | 0.77 | 0.85 | < 1 | 0.74 | 0.7 | 0.76 | < 0.1 | 1 | 0.7516 | 0.77 | 0.7594 | 0.79 | 0.88 | 0.79 | 0.78 |  |  | 23
| 1 | 3 | 0.91 | 0.81 | 0.29 | 0.87486 | 0.6 | 1 | 0.76 | 0.83 | 1.1 | 0.8 | 0.79 | 0.76 | 0.77 | 0.79 | < 1 | 0.73 | 0.7 | 0.74 | < 0.1 | 1 | 0.74215 | 0.81 | 0.7634 | 0.78 | 0.86 | 0.81 | 0.76 |  |  | 16
| 1 | 4 | 0.92 | 0.69 | 0.27 | 0.85832 | 0.7 | 1 | 0.76 | 0.78 | 1.2 | 0.9 | 0.8 | 0.8 | 0.74 | 0.77 | < 1 | 0.72 | 0.7 | 0.71 | < 0.1 | 1 | 0.74369 | 0.77 | 0.746 | 0.74 | 0.85 | 0.76 | 0.8 |  |  | 0.77953
| 1 | 5 | 0.96 | 0.84 | 0.3 | 0.85528 | 0.7 | 1 | 0.93 | 0.82 | 1.1 | 0.9 | 0.8 | 0.81 | 0.77 | 0.8 | < 1 | 0.79 | 0.7 | 0.76 | < 0.1 | 1 | 0.73565 | 0.78 | 0.7223 | 0.77 | 0.92 | 0.81 | 0.83 |  |  | 20
| 1 | 6 | 0.94 | 0.82 | 0.3 | 0.89348 | 0.7 | 0.9 | 0.88 | 0.8 | 1.2 | 1 | 0.79 | 0.86 | 0.78 | 0.76 | < 1 | 0.83 | 0.7 | 0.73 | < 0.1 | 1 | 0.74411 | 0.8 | 0.7616 | 0.74 | 0.98 | 0.76 | 0.83
| Mean |  | 0.925 | 0.77833 | 0.28833 | 0.87541 | 0.65 | 1.0167 | 0.815 | 0.80833 | 1.1667 | 0.88333 | 0.79167 | 0.81 | 0.76167 | 0.79333 |  | 0.77167 | 0.7 | 0.74333 |  | 1 | 0.74301 | 0.78333 | 0.75195 | 0.76833 | 0.89167 | 0.78667 | 0.80333
| Median |  | 0.92 | 0.795 | 0.29 | 0.86746 | 0.65 | 1 | 0.78 | 0.81 | 1.2 | 0.9 | 0.79 | 0.805 | 0.77 | 0.79 |  | 0.765 | 0.7 | 0.75 |  | 1 | 0.74292 | 0.775 | 0.7592 | 0.775 | 0.87 | 0.79 | 0.81
| Std Dev. |  | 0.021679 | 0.057764 | 0.01169 | 0.022311 | 0.054772 | 0.075277 | 0.072042 | 0.036009 | 0.05164 | 0.075277 | 0.0075277 | 0.033466 | 0.017224 | 0.031411 |  | 0.047924 | 1.2162e-16 | 0.020656 |  | 0 | 0.0051944 | 0.017512 | 0.015773 | 0.023166 | 0.049967 | 0.022509 | 0.028752
| Rel.Std.Dev. |  | 2.344% | 7.421% | 4.054% | 2.549% | 8.427% | 7.404% | 8.839% | 4.455% | 4.426% | 8.522% | 0.9509% | 4.132% | 2.261% | 3.959% |  | 6.21% | 1.737e-14% | 2.779% |  | 0% | 0.6991% | 2.236% | 2.098% | 3.015% | 5.604% | 2.861% | 3.579%
| PDM3 |  | 18.66% | -0.153% | -63.01% | 12.3% | -16.62% | 30.42% | 4.551% | 3.695% | 49.66% | 13.32% | 1.557% | 3.909% | -2.291% | 1.771% |  | -1.008% | -10.2% | -4.643% |  | 28.28% | -4.684% | 0.4884% | -3.538% | -1.436% | 14.39% | 0.916% | 3.054%
| Z-Score (Absolute) |  | 3.58 | 0.13 | 12.52 | 2.33 | NA | NA | 0.8 | 0.63 | NA | NA | 0.21 | 0.67 | 0.55 | 0.25 | 7.16 | 0.3 | NA | 1.01 | 18.54 | NA | 1.02 | 0 | 0.79 | 0.38 | 2.74 | 0.08 | 0.51
Results from laboratories 1.07, 1.08, 1.12, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 241b (Certified Value 14.8 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 15.35 | 15 | 12.97 | 13.2149 | 15.1095 | 13.9 | 18 | 14.01 | 15 | 14.7 | 16.4 | 14.35 | 14.5 | < 10 | 13 | 14.63 | 15 | 14.5124 | 15.4 | 7.9558 | 17.5 | 14.75 | 15.35 | 16.3 | [5 empty] | 1
| 1 | 2 | 15.53 | 14.84 | 13.46 | 13.4771 | 15.8201 | 13.5 | 17 | 14.49 | 15 | 14.75 | 15.05 | 14.65 | 14.2 | < 10 | 14 | 14.12 | 15 | 14.7317 | 14 | 7.8935 | 17.3 | 14.84 | 15.17 | 15.9 | [5 empty] | 24
| 1 | 3 | 16.74 | 14.65 | 13.49 | 13.3028 | 15.1685 | 13.6 | 18 | 14.61 | 15 | 14.55 | 14.95 | 13.95 | 13.4 | < 10 | 13 | 13.75 | 15.5 | 14.9866 | 14.95 | 7.9194 | 17.4 | 14.61 | 15.42 | 15.6 | [5 empty] | 16
| 1 | 4 | 13.58 | 14.59 | 13.08 | 13.1991 | 15.2333 | 13.7 | 17 | 14.16 | 15 | 14.6 | 16 | 13.95 | 13.7 | < 10 | 13 | 14.51 | 14.5 | 14.8969 | 14.8 | 7.8555 | 17.4 | 14.48 | 15.58 | 15.9 | [5 empty] | 14.8499
| 1 | 5 | 16.88 | 14.85 | 13.43 | 13.5414 | 15.3478 | 13.5 | 17 | 14.29 | 15 | 14.6 | 16.7 | 14.25 | 13.85 | < 10 | 14 | 13.98 | 15 | 14.784 | 15.35 | 7.5903 | 17.4 | 15.18 | 16.12 | 16.6 | [5 empty] | 21
| 1 | 6 | 15.96 | 14.76 | 13.25 | 13.3835 | 16.098 | 14.2 | 18 | 14.58 | 15 | 14.55 | 16.3 | 13.95 | 13.6 | 12 | 14 | 14.1 | 14.5 | 15.0006 | 15 | 7.8775 | 17.2 | 15.28 | 15.66 | 17.6
| Mean |  | 15.6733 | 14.7817 | 13.28 | 13.3531 | 15.4629 | 13.7333 | 17.5 | 14.3567 | 15 | 14.625 | 15.9 | 14.1833 | 13.875 | 12 | 13.5 | 14.1817 | 14.9167 | 14.8187 | 14.9167 | 7.8487 | 17.3667 | 14.8567 | 15.55 | 16.3167
| Median |  | 15.745 | 14.8 | 13.34 | 13.3432 | 15.2906 | 13.65 | 17.5 | 14.39 | 15 | 14.6 | 16.15 | 14.1 | 13.775 | 12 | 13.5 | 14.11 | 15 | 14.8405 | 14.975 | 7.8855 | 17.4 | 14.795 | 15.5 | 16.1
| Std Dev. |  | 1.1986 | 0.14851 | 0.21726 | 0.13934 | 0.40209 | 0.27325 | 0.54772 | 0.24295 | 0 | 0.082158 | 0.7328 | 0.28752 | 0.40712 |  | 0.54772 | 0.33054 | 0.37639 | 0.18439 | 0.50662 | 0.13124 | 0.10328 | 0.31564 | 0.32839 | 0.71949
| Rel.Std.Dev. |  | 7.647% | 1.005% | 1.636% | 1.044% | 2.6% | 1.99% | 3.13% | 1.692% | 0% | 0.5618% | 4.609% | 2.027% | 2.934% |  | 4.057% | 2.331% | 2.523% | 1.244% | 3.396% | 1.672% | 0.5947% | 2.125% | 2.112% | 4.41%
| PDM3 |  | 5.545% | -0.4595% | -10.57% | -10.08% | 4.128% | -7.519% | 17.85% | -3.321% | 1.011% | -1.514% | 7.071% | -4.489% | -6.565% | -19.19% | -9.09% | -4.5% | 0.4496% | -0.2101% | 0.4496% | -47.15% | 16.95% | 0.04559% | 4.715% | 9.877%
| Z-Score (Absolute) |  | 0.83 | 0.02 | 1.45 | 1.38 | 0.63 | 1.02 | NA | 0.42 | NA | 0.17 | 1.05 | 0.59 | 0.88 | NA | NA | 0.59 | 0.11 | 0.02 | 0.11 | 6.64 | 2.45 | 0.05 | 0.72 | 1.45
Results from laboratories 1.08, 1.13, 1.19 and 1.23 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 241b (Certified Value 39.8 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 41 | 39.8 | 39.54 | 30.9801 | 39.869 | 40 | 55.9 | 39 | 38.8 | 31 | 38.6 | 40.8 | 43.7 | 40.5 | 39.8 | 39 | 40.1 | 37.3 | 42.6 | 33.6 | 35 | 40.4669 | 39.6 | 34.8376 | 47.1 | 41.4 | 36.5 | 40.8 |  | 1
| 1 | 2 | 41.5 | 40.8 | 39.02 | 32.488 | 41.3923 | 40 | 54.5 | 40 | 41.4 | 31 | 39.9 | 40.7 | 42.5 | 41 | 40.8 | 37 | 40.5 | 37.9 | 41 | 33.89 | 35 | 40.3098 | 35.3 | 35.1845 | 46.4 | 40.6 | 37.4 | 39.4 |  | 25
| 1 | 3 | 41.8 | 40.3 | 39.18 | 32.1149 | 39.5158 | 40 | 49.5 | 41 | 40.5 | 31 | 39.3 | 40.8 | 41.9 | 39.4 | 39.5 | 38 | 43.8 | 36.8 | 41.9 | 34.56 | 35 | 40.8876 | 40.7 | 35.1392 | 46.3 | 40.9 | 36.2 | 38.7 |  | 16
| 1 | 4 | 42.2 | 40 | 39.35 | 31.4463 | 39.9266 | 40 | 50.1 | 40 | 40.3 | 31 | 39.2 | 41.4 | 44.5 | 39.5 | 39.7 | 40 | 40.3 | 38 | 41.6 | 33.04 | 35 | 41.3759 | 38.7 | 36.709 | 46 | 40.5 | 35.5 | 39.3 |  | 39.7593
| 1 | 5 | 45.4 | 41.1 | 39.12 | 32.4217 | 39.7379 | 41 | 49.3 | 41 | 40.5 | 31 | 38.5 | 40.9 | 45.6 | 41.2 | 40.8 | 39 | 43.4 | 38.9 | 42.4 | 33.55 | 35 | 40.2893 | 40.6 | 34.3067 | 47.4 | 41.6 | 36.9 | 40.7 |  | 22
| 1 | 6 | 42.5 | 39.6 | 39.74 | 31.611 | 40.757 | 40 | 45.6 | 42 | 41.1 | 31 | 39.5 | 41 | 45.1 | 39.8 | 40.8 | 40 | 42.5 | 37 | 42.2 | 34.66 | 35 | 40.0788 | 37.2 | 34.3856 | 47.1 | 42.4 | 36.1 | 43.2
| Mean |  | 42.4 | 40.2667 | 39.325 | 31.8437 | 40.1998 | 40.1667 | 50.8167 | 40.5 | 40.4333 | 31 | 39.1667 | 40.9333 | 43.8833 | 40.2333 | 40.2333 | 38.8333 | 41.7667 | 37.65 | 41.95 | 33.8833 | 35 | 40.568 | 38.6833 | 35.0938 | 46.7167 | 41.2333 | 36.4333 | 40.35
| Median |  | 42 | 40.15 | 39.265 | 31.863 | 39.8978 | 40 | 49.8 | 40.5 | 40.5 | 31 | 39.25 | 40.85 | 44.1 | 40.15 | 40.3 | 39 | 41.5 | 37.6 | 42.05 | 33.745 | 35 | 40.3884 | 39.15 | 34.9884 | 46.75 | 41.15 | 36.35 | 40.05
| Std Dev. |  | 1.5608 | 0.58538 | 0.27392 | 0.59679 | 0.72079 | 0.40825 | 3.7728 | 1.0488 | 0.90259 | 0 | 0.53541 | 0.25033 | 1.4621 | 0.77632 | 0.62823 | 1.169 | 1.6657 | 0.77653 | 0.58566 | 0.6267 | 0 | 0.47933 | 2.1085 | 0.87245 | 0.55648 | 0.7174 | 0.66232 | 1.6233
| Rel.Std.Dev. |  | 3.681% | 1.454% | 0.6965% | 1.874% | 1.793% | 1.016% | 7.424% | 2.59% | 2.232% | 0% | 1.367% | 0.6116% | 3.332% | 1.93% | 1.561% | 3.01% | 3.988% | 2.062% | 1.396% | 1.85% | 0% | 1.182% | 5.451% | 2.486% | 1.191% | 1.74% | 1.818% | 4.023%
| PDM3 |  | 6.642% | 1.276% | -1.092% | -19.91% | 1.108% | 1.024% | 27.81% | 1.863% | 1.695% | -22.03% | -1.491% | 2.953% | 10.37% | 1.192% | 1.192% | -2.329% | 5.049% | -5.305% | 5.51% | -14.78% | -11.97% | 2.034% | -2.706% | -11.73% | 17.5% | 3.707% | -8.365% | 1.486%
| Z-Score (Absolute) |  | 1 | 0.02 | 0.42 | 3.86 | 0.02 | 0.03 | 4.87 | 0.12 | 0.09 | 4.25 | 0.49 | 0.32 | 1.68 | 0 | 0 | 0.64 | 0.7 | 1.19 | 0.79 | 2.92 | 2.41 | 0.15 | 0.71 | 2.36 | 2.98 | 0.46 | 1.75 | 0.05
| Analytical results for Cr in OREAS 241b (Certified Value 119 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 100 | 130 | 110.889 | 120 | 168 | 114 | 95 | 119 | 146 | 106 | 131 | 111 | 106 | 119 | 110 | 116 | 122 | 116.24 | 131.301 | 130 | 138.5425 | 113 | 122.0824 | 117 | 123 | 133 | 116 |  |  | 1
| 1 | 2 | 97 | 141 | 116.5851 | 121 | 169 | 111 | 103 | 120 | 140 | 108 | 129 | 112 | 110 | 113 | 106 | 113 | 122 | 121.45 | 129.816 | 130 | 137.9436 | 112 | 121.4839 | 130 | 123 | 129 | 114 |  |  | 26
| 1 | 3 | 94 | 140 | 109.7524 | 100 | 171 | 114 | 124 | 119 | 140 | 114 | 127 | 111 | 109 | 116 | 111 | 115 | 117 | 114.41 | 132.255 | 135 | 140.5505 | 114 | 125.1783 | 124 | 124 | 130 | 118 |  |  | 16
| 1 | 4 | 99 | 139 | 113.2723 | 86 | 163 | 113 | 126 | 122 | 137 | 114 | 127 | 112 | 109 | 122 | 100 | 117 | 123 | 112.79 | 141.399 | 135 | 140.026 | 114 | 121.3872 | 127 | 121 | 129 | 109 |  |  | 119.3824
| 1 | 5 | 97 | 141 | 112.8937 | 81 | 167 | 112 | 129 | 122 | 136 | 107 | 122 | 123 | 111 | 119 | 104 | 119 | 121 | 107.27 | 134.208 | 135 | 138.067 | 114 | 124.1802 | 121 | 124 | 128 | 98.8 |  |  | 23
| 1 | 6 | 99 | 136 | 116.3507 | 90 | 170 | 110 | 119 | 119 | 138 | 107 | 123 | 113 | 111 | 121 | 102 | 114 | 120 | 99.92 | 127.512 | 140 | 138.702 | 112 | 121.1845 | 134 | 126 | 132 | 112
| Mean |  | 97.6667 | 137.8333 | 113.2905 | 99.6667 | 168 | 112.3333 | 116 | 120.1667 | 139.5 | 109.3333 | 126.5 | 113.6667 | 109.3333 | 118.3333 | 105.5 | 115.6667 | 120.8333 | 112.0133 | 132.7485 | 134.1667 | 138.9719 | 113.1667 | 122.5828 | 125.5 | 123.5 | 130.1667 | 111.3
| Median |  | 98 | 139.5 | 113.083 | 95 | 168.5 | 112.5 | 121.5 | 119.5 | 139 | 107.5 | 127 | 112 | 109.5 | 119 | 105 | 115.5 | 121.5 | 113.6 | 131.778 | 135 | 138.6223 | 113.5 | 121.7832 | 125.5 | 123.5 | 129.5 | 113
| Std Dev. |  | 2.1602 | 4.2622 | 2.7807 | 17.3051 | 2.8284 | 1.633 | 13.7986 | 1.472 | 3.5637 | 3.6697 | 3.4496 | 4.6332 | 1.8619 | 3.3267 | 4.3704 | 2.1602 | 2.137 | 7.5101 | 4.8025 | 3.7639 | 1.0711 | 0.98319 | 1.6811 | 6.1563 | 1.6432 | 1.9408 | 6.8746
| Rel.Std.Dev. |  | 2.212% | 3.092% | 2.455% | 17.36% | 1.684% | 1.454% | 11.9% | 1.225% | 2.555% | 3.356% | 2.727% | 4.076% | 1.703% | 2.811% | 4.143% | 1.868% | 1.769% | 6.705% | 3.618% | 2.805% | 0.7708% | 0.8688% | 1.371% | 4.905% | 1.331% | 1.491% | 6.177%
| PDM3 |  | -18.19% | 15.46% | -5.103% | -16.51% | 40.72% | -5.905% | -2.833% | 0.6569% | 16.85% | -8.418% | 5.962% | -4.788% | -8.418% | -0.8788% | -11.63% | -3.113% | 1.215% | -6.173% | 11.2% | 12.38% | 16.41% | -5.207% | 2.681% | 5.124% | 3.449% | 9.033% | -6.77%
| Z-Score (Absolute) |  | 1.94 | 1.83 | 0.47 | 1.76 | 4.67 | 0.56 | 0.22 | 0.17 | 1.99 | 0.85 | 0.77 | 0.44 | 0.85 | 0 | 1.21 | 0.25 | 0.24 | 0.59 | 1.36 | 1.49 | 1.94 | 0.49 | 0.4 | 0.67 | 0.49 | 1.11 | 0.66
| Analytical results for Cs in OREAS 241b (Certified Value 1.22 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | < 5 | 1.2 | 1.0878 | 1.11 | 1.6 | 1.2 | 1.18 | 1.4 | 1.19 | 1.27 | 1.25 | 1.2 | 1.2 | 1.16 | 1.5 | 1.1727 | 1.27 | 1.159 | 1.14 | 1.35 | 1.24 | 1.26 | [7 empty] | 1
| 1 | 2 | < 5 | 1.2 | 1.1493 | 1.12 | 1.5 | 1.2 | 1.26 | 1.4 | 1.2 | 1.24 | 1.23 | 1.22 | 1.5 | 1.2 | 1.5 | 1.1165 | 1.2 | 1.1536 | 1.14 | 1.3 | 1.25 | 1.24 | [7 empty] | 27
| 1 | 3 | < 5 | 1.21 | 1.0968 | 1.1 | 1.4 | 1.2 | 1.24 | 1.3 | 1.22 | 1.21 | 1.2 | 1.17 | 1.4 | 1.22 | 1.5 | 1.1892 | 1.27 | 1.1413 | 1.13 | 1.34 | 1.26 | 1.21 | [7 empty] | 16
| 1 | 4 | < 5 | 1.17 | 1.0802 | 1.12 | 1.4 | 1.1 | 1.22 | 1.4 | 1.2 | 1.28 | 1.2 | 1.16 | 1.2 | 1.14 | 1.5 | 1.1448 | 1.23 | 1.1681 | 1.08 | 1.36 | 1.24 | 1.23 | [7 empty] | 1.2178
| 1 | 5 | < 5 | 1.25 | 1.1438 | 1.09 | 1.4 | 1.2 | 1.27 | 1.3 | 1.16 | 1.32 | 1.24 | 1.17 | 1.4 | 1.19 | 1 | 1.1981 | 1.28 | 1.1285 | 1.14 | 1.37 | 1.27 | 1.28 | [7 empty] | 24
| 1 | 6 | < 5 | 1.19 | 1.1054 | 1.13 | 1.3 | 1.2 | 1.26 | 1.4 | 1.2 | 1.32 | 1.24 | 1.14 | 1.3 | 1.18 | 1.5 | 1.1897 | 1.26 | 1.1305 | 1.16 | 1.46 | 1.24 | 1.35
| Mean |  |  | 1.2033 | 1.1106 | 1.1117 | 1.4333 | 1.1833 | 1.2383 | 1.3667 | 1.195 | 1.2733 | 1.2267 | 1.1767 | 1.3333 | 1.1817 | 1.4167 | 1.1685 | 1.2517 | 1.1468 | 1.1317 | 1.3633 | 1.25 | 1.2617
| Median |  |  | 1.2 | 1.1011 | 1.115 | 1.4 | 1.2 | 1.25 | 1.4 | 1.2 | 1.275 | 1.235 | 1.17 | 1.35 | 1.185 | 1.5 | 1.1809 | 1.265 | 1.1475 | 1.14 | 1.355 | 1.245 | 1.25
| Std Dev. |  |  | 0.026583 | 0.029193 | 0.01472 | 0.10328 | 0.040825 | 0.033714 | 0.05164 | 0.019748 | 0.043665 | 0.021602 | 0.028752 | 0.12111 | 0.028577 | 0.20412 | 0.03175 | 0.030605 | 0.015997 | 0.027142 | 0.053166 | 0.012649 | 0.049565
| Rel.Std.Dev. |  |  | 2.209% | 2.629% | 1.324% | 7.206% | 3.45% | 2.723% | 3.779% | 1.653% | 3.429% | 1.761% | 2.443% | 9.083% | 2.418% | 14.41% | 2.717% | 2.445% | 1.395% | 2.398% | 3.9% | 1.012% | 3.929%
| PDM3 |  |  | -1.191% | -8.808% | -8.718% | 17.7% | -2.833% | 1.683% | 12.22% | -1.875% | 4.557% | 0.7251% | -3.381% | 9.484% | -2.97% | 16.33% | -4.052% | 2.778% | -5.83% | -7.076% | 11.95% | 2.641% | 3.599%
| Z-Score (Absolute) |  | 14.26 | 0.33 | 1.37 | 1.36 | 2.26 | 0.55 | 0.07 | 1.51 | 0.42 | 0.46 | 0.07 | 0.63 | 1.13 | 0.57 | 2.07 | 0.72 | 0.22 | 0.96 | 1.13 | 1.47 | 0.2 | 0.33
| Analytical results for Cu in OREAS 241b (Certified Value 166 ppm) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 166.6 | 163.4 | 156.74 | 171.43 | 166.0073 | 178 | 219 | 161 | 157.7 | 160 | 169.9 | 165.5 | 183.5 | 168.5 | 170.5 | 170 | 169 | 150.8 | 172 | 34.57 | 167.499 | 165 | 158.2809 | 164 | 109.4548 | 179 | 162.2 | 167.4 | 161 | 1
| 1 | 2 | 169.3 | 166.6 | 165.89 | 177.84 | 171.7165 | 171 | 210 | 162 | 165 | 163 | 169 | 166 | 177 | 166 | 170 | 163 | 168 | 147.8 | 175 | 35.67 | 165.069 | 160 | 155.7144 | 160 | 111.3398 | 175 | 164 | 167.6 | 158 | 28
| 1 | 3 | 172.7 | 164.9 | 154.32 | 175.92 | 165.2982 | 170 | 192 | 165 | 162.3 | 160 | 168.7 | 170 | 172 | 165.5 | 170.5 | 167 | 168 | 148.6 | 171 | 33.08 | 165.627 | 165 | 159.526 | 166.5 | 110.4211 | 181 | 162.4 | 168.7 | 154 | 16
| 1 | 4 | 174.3 | 162.8 | 154.68 | 175.32 | 166.8763 | 173 | 193 | 159 | 159.5 | 162 | 165.9 | 171.5 | 179.5 | 163.5 | 170 | 175 | 166 | 148 | 173 | 34.51 | 164.889 | 160 | 157.3272 | 164 | 113.6061 | 181 | 169 | 167.9 | 160 | 166.3441
| 1 | 5 | 173.6 | 164.2 | 157.97 | 172.45 | 166.9102 | 177 | 192 | 162 | 164.3 | 162 | 166.2 | 164.5 | 183.5 | 179.5 | 171 | 176 | 166 | 152.3 | 172 | 41.79 | 161.658 | 165 | 155.4422 | 168 | 108.3952 | 178 | 164.8 | 165.5 | 155 | 25
| 1 | 6 | 172.1 | 161.2 | 154.31 | 178.78 | 170.3089 | 174 | 178 | 164 | 164.2 | 162 | 169.5 | 166 | 179 | 165.5 | 171 | 178 | 165 | 148.1 | 176 | 38.59 | 162.414 | 165 | 155.8578 | 164.5 | 105.5275 | 183 | 171.4 | 169.7 | 155
| Mean |  | 171.4333 | 163.85 | 157.3183 | 175.29 | 167.8529 | 173.8333 | 197.3333 | 162.1667 | 162.1667 | 161.5 | 168.2 | 167.25 | 179.0833 | 168.0833 | 170.5 | 171.5 | 167 | 149.2667 | 173.1667 | 36.3683 | 164.526 | 163.3333 | 157.0248 | 164.5 | 109.7908 | 179.5 | 165.6333 | 167.8 | 157.1667
| Median |  | 172.4 | 163.8 | 155.71 | 175.62 | 166.8932 | 173.5 | 192.5 | 162 | 163.25 | 162 | 168.85 | 166 | 179.25 | 165.75 | 170.5 | 172.5 | 167 | 148.35 | 172.5 | 35.12 | 164.979 | 165 | 156.5925 | 164.25 | 109.938 | 180 | 164.4 | 167.75 | 156.5
| Std Dev. |  | 2.9283 | 1.8501 | 4.455 | 2.8998 | 2.5587 | 3.1885 | 14.6924 | 2.137 | 2.9595 | 1.2247 | 1.7181 | 2.8062 | 4.3292 | 5.8174 | 0.44721 | 5.8224 | 1.5492 | 1.85 | 1.9408 | 3.2348 | 2.1524 | 2.582 | 1.6437 | 2.7203 | 2.7419 | 2.8107 | 3.7468 | 1.4085 | 2.9269
| Rel.Std.Dev. |  | 1.708% | 1.129% | 2.832% | 1.654% | 1.524% | 1.834% | 7.445% | 1.318% | 1.825% | 0.7584% | 1.021% | 1.678% | 2.417% | 3.461% | 0.2623% | 3.395% | 0.9277% | 1.239% | 1.121% | 8.894% | 1.308% | 1.581% | 1.047% | 1.654% | 2.497% | 1.566% | 2.262% | 0.8394% | 1.862%
| PDM3 |  | 3.059% | -1.499% | -5.426% | 5.378% | 0.907% | 4.502% | 18.63% | -2.511% | -2.511% | -2.912% | 1.116% | 0.5446% | 7.658% | 1.046% | 2.498% | 3.1% | 0.3943% | -10.27% | 4.101% | -78.14% | -1.093% | -1.81% | -5.602% | -1.109% | -34% | 7.909% | -0.4273% | 0.8752% | -5.517%
| Z-Score (Absolute) |  | 0.62 | 0.44 | 1.35 | 1.16 | 0.12 | 0.95 | 4.23 | 0.67 | 0.67 | 0.77 | 0.17 | 0.03 | 1.69 | 0.15 | 0.49 | 0.63 | 0 | 2.47 | 0.86 | 18.22 | 0.35 | 0.51 | 1.39 | 0.35 | 7.98 | 1.74 | 0.19 | 0.11 | 1.37
| Analytical results for Dy in OREAS 241b (Certified Value 3.63 ppm) | [30 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.65 | 3.27 | 3.2756 | 3.6 | 4.37 | 3.65 | 3.9 | 3.6835 | [21 empty] | 1
| 1 | 2 | 3.72 | 3.31 | 3.2967 | 3.6 | 4.24 | 3.6 | 3.9 | 3.722 | [21 empty] | 29
| 1 | 3 | 3.69 | 3.26 | 3.3255 | 3.5 | 3.95 | 3.61 | 4.2 | 3.9304 | [21 empty] | 16
| 1 | 4 | 3.6 | 3.25 | 3.2813 | 3.7 | 3.91 | 3.63 | 3.9 | 3.7882 | [21 empty] | 3.6304
| 1 | 5 | 3.55 | 3.3 | 3.2692 | 3.6 | 3.88 | 3.55 | 3.7 | 3.759 | [21 empty] | 26
| 1 | 6 | 3.57 | 3.24 | 3.2859 | 3.6 | 3.7 | 3.57 | 3.8 | 3.7591
| Mean |  | 3.63 | 3.2717 | 3.289 | 3.6 | 4.0083 | 3.6017 | 3.9 | 3.7737
| Median |  | 3.625 | 3.265 | 3.2836 | 3.6 | 3.93 | 3.605 | 3.9 | 3.7591
| Std Dev. |  | 0.067823 | 0.027869 | 0.020166 | 0.063246 | 0.24863 | 0.037103 | 0.16733 | 0.08486
| Rel.Std.Dev. |  | 1.868% | 0.8518% | 0.6131% | 1.757% | 6.203% | 1.03% | 4.291% | 2.249%
| PDM3 |  | -0.01053% | -9.881% | -9.403% | -0.8369% | 10.41% | -0.791% | 7.427% | 3.948%
| Z-Score (Absolute) |  | 0.04 | 1.05 | 1 | 0.05 | 1.2 | 0.04 | 0.87 | 0.48
| Analytical results for Er in OREAS 241b (Certified Value 2.21 ppm) | [30 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.21 | 1.96 | 1.9595 | 2.1 | 2.67 | 2.28 | 2.6 | 2.3566 | [21 empty] | 1
| 1 | 2 | 2.27 | 1.97 | 2.0389 | 2 | 2.59 | 2.15 | 2.4 | 2.3786 | [21 empty] | 30
| 1 | 3 | 2.19 | 1.94 | 1.949 | 2.1 | 2.4 | 2.1 | 2.6 | 2.4634 | [21 empty] | 16
| 1 | 4 | 2.22 | 1.92 | 2.0712 | 2.1 | 2.39 | 2.17 | 2.4 | 2.3558 | [21 empty] | 2.209
| 1 | 5 | 2.2 | 1.99 | 2.0845 | 2 | 2.43 | 2.14 | 2.4 | 2.3721 | [21 empty] | 27
| 1 | 6 | 2.23 | 1.94 | 2.0157 | 2 | 2.27 | 2.13 | 2.3 | 2.3689
| Mean |  | 2.22 | 1.9533 | 2.0198 | 2.05 | 2.4583 | 2.1617 | 2.45 | 2.3826
| Median |  | 2.215 | 1.95 | 2.0273 | 2.05 | 2.415 | 2.145 | 2.4 | 2.3705
| Std Dev. |  | 0.028284 | 0.025033 | 0.056317 | 0.054772 | 0.14593 | 0.062423 | 0.12247 | 0.040588
| Rel.Std.Dev. |  | 1.274% | 1.282% | 2.788% | 2.672% | 5.936% | 2.888% | 4.999% | 1.704%
| PDM3 |  | 0.4978% | -11.57% | -8.566% | -7.198% | 11.29% | -2.143% | 10.91% | 7.857%
| Z-Score (Absolute) |  | 0.11 | 0.88 | 0.64 | 0.52 | 0.99 | 0.11 | 0.96 | 0.71
| Analytical results for Eu in OREAS 241b (Certified Value 0.97 ppm) | [30 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
2
| 1 | 1 | 0.9 | 0.84 | 1.1818 | 0.85 | 1.12 | 0.99 | 1.1 | 0.90287 | [21 empty] | 1
| 1 | 2 | 0.93 | 0.86 | 1.2317 | 0.83 | 1.07 | 0.99 | 1 | 0.9412 | [21 empty] | 31
| 1 | 3 | 0.91 | 0.83 | 1.2048 | 0.84 | 0.98 | 0.97 | 1.1 | 0.95334 | [21 empty] | 16
| 1 | 4 | 0.9 | 0.81 | 1.218 | 0.88 | 0.98 | 0.95 | 1.1 | 0.96579 | [21 empty] | 0.9703
| 1 | 5 | 0.88 | 0.83 | 1.2272 | 0.87 | 0.98 | 0.93 | 1 | 0.93152 | [21 empty] | 28
| 1 | 6 | 0.89 | 0.82 | 1.1793 | 0.86 | 0.93 | 0.98 | 1 | 0.9371
| Mean |  | 0.90167 | 0.83167 | 1.2071 | 0.855 | 1.01 | 0.96833 | 1.05 | 0.93864
| Median |  | 0.9 | 0.83 | 1.2114 | 0.855 | 0.98 | 0.975 | 1.05 | 0.93915
| Std Dev. |  | 0.017224 | 0.017224 | 0.022579 | 0.018708 | 0.070427 | 0.024014 | 0.054772 | 0.021412
| Rel.Std.Dev. |  | 1.91% | 2.071% | 1.87% | 2.188% | 6.973% | 2.48% | 5.216% | 2.281%
| PDM3 |  | -7.074% | -14.29% | 24.41% | -11.88% | 4.091% | -0.2031% | 8.214% | -3.264%
| Z-Score (Absolute) |  | 0.46 | 1.07 | 2.24 | 0.87 | 0.5 | 0.13 | 0.85 | 0.13
| Analytical results for Fe in OREAS 241b (Certified Value 7.54 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 7.56 | 7.65 | 7.6689 | 7.536 | 6.87 | 8.323 | 7.51 | 7.3 | 7.28 | 7.31 | 7.55 | 8.22 | 7.61 | 7.72 | 7.12 | 8.07 | 6.93 | 7.54 | 7.62 | 7.4278 | 7.745 | 7.54 | 7.5723 | 7.27 | 7.12 | 7.99 | 7.2 |  |  | 1
| 1 | 2 | 7.66 | 7.66 | 7.6821 | 7.6438 | 6.82 | 8.463 | 7.7 | 7.66 | 7.25 | 7.39 | 7.51 | 7.85 | 7.51 | 7.56 | 6.8 | 7.9 | 6.86 | 7.7 | 7.09 | 7.2947 | 7.65 | 7.2 | 7.5393 | 7.17 | 7.05 | 7.95 | 7.5 |  |  | #N/A
| 1 | 3 | 7.7 | 7.62 | 7.6019 | 7.5059 | 6.82 | 8.463 | 7.87 | 7.47 | 7.22 | 7.36 | 7.85 | 7.79 | 7.44 | 7.7 | 7.02 | 8.01 | 6.94 | 7.47 | 7.62 | 7.3105 | 7.725 | 7.63 | 7.7563 | 7.54 | 7.08 | 8.1 | 7.1 |  |  | 16
| 1 | 4 | 7.76 | 7.71 | 7.6115 | 7.5775 | 6.76 | 7.974 | 7.67 | 7.4 | 7.29 | 7.26 | 7.87 | 8.03 | 7.38 | 7.56 | 7.34 | 7.84 | 6.98 | 7.55 | 6.98 | 7.2996 | 7.72 | 7.54 | 7.5203 | 7.39 | 6.95 | 7.98 | 7.2 |  |  | 7.5421
| 1 | 5 | 7.76 | 7.72 | 7.6281 | 7.5941 | 6.76 | 8.253 | 7.78 | 7.46 | 7.24 | 7.26 | 7.62 | 8.18 | 8.24 | 7.65 | 7.16 | 7.95 | 7.11 | 7.44 | 6.75 | 7.2959 | 7.735 | 7.69 | 7.8033 | 7.48 | 7.16 | 7.96 | 7.8 |  |  | 29
| 1 | 6 | 7.71 | 7.51 | 7.6269 | 7.6294 | 7.32 | 8.323 | 7.92 | 7.55 | 7.29 | 7.33 | 7.64 | 7.91 | 7.51 | 7.75 | 7.2 | 7.76 | 6.91 | 7.78 | 6.83 | 7.3206 | 7.745 | 7.38 | 7.4683 | 7.55 | 7.24 | 8.04 | 7.8
| Mean |  | 7.6917 | 7.645 | 7.6366 | 7.5811 | 6.8917 | 8.2998 | 7.7417 | 7.4733 | 7.2617 | 7.3183 | 7.6733 | 7.9967 | 7.615 | 7.6567 | 7.1067 | 7.9217 | 6.955 | 7.58 | 7.1483 | 7.3249 | 7.72 | 7.4967 | 7.61 | 7.4 | 7.1 | 8.0033 | 7.4333
| Median |  | 7.705 | 7.655 | 7.6275 | 7.5858 | 6.82 | 8.323 | 7.74 | 7.465 | 7.265 | 7.32 | 7.63 | 7.97 | 7.51 | 7.675 | 7.14 | 7.925 | 6.935 | 7.545 | 7.035 | 7.3051 | 7.73 | 7.54 | 7.5558 | 7.435 | 7.1 | 7.985 | 7.35
| Std Dev. |  | 0.074944 | 0.076092 | 0.031944 | 0.053146 | 0.21395 | 0.18038 | 0.14851 | 0.12356 | 0.029269 | 0.052694 | 0.15214 | 0.17682 | 0.31577 | 0.08165 | 0.18316 | 0.11303 | 0.085499 | 0.13312 | 0.38384 | 0.051395 | 0.035777 | 0.17918 | 0.1366 | 0.15388 | 0.098995 | 0.056804 | 0.31411
| Rel.Std.Dev. |  | 0.9744% | 0.9953% | 0.4183% | 0.701% | 3.105% | 2.173% | 1.918% | 1.653% | 0.4031% | 0.72% | 1.983% | 2.211% | 4.147% | 1.066% | 2.577% | 1.427% | 1.229% | 1.756% | 5.37% | 0.7017% | 0.4634% | 2.39% | 1.795% | 2.08% | 1.394% | 0.7098% | 4.226%
| PDM3 |  | 1.983% | 1.364% | 1.252% | 0.517% | -8.624% | 10.05% | 2.646% | -0.912% | -3.718% | -2.967% | 1.74% | 6.027% | 0.9663% | 1.519% | -5.774% | 5.032% | -7.785% | 0.5023% | -5.221% | -2.881% | 2.358% | -0.6027% | 0.9001% | -1.884% | -5.862% | 6.115% | -1.442%
| Z-Score (Absolute) |  | 0.46 | 0.27 | 0.23 | 0 | 2.9 | 3.02 | 0.67 | 0.45 | 1.34 | 1.1 | 0.39 | 1.75 | 0.14 | 0.32 | 1.99 | 1.43 | 2.63 | 0 | 1.82 | 1.08 | 0.58 | 0.35 | 0.12 | 0.76 | 2.02 | 1.77 | 0.62
| Analytical results for Ga in OREAS 241b (Certified Value 15.3 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 16.7 | 14.94 | 13.6228 | 13.7518 | 12.9 | 18 | 15.2 | 15.14 | 18 | 15 | 15 | 17.1 | 15.15 | 14.95 | 10 | 15.3 | 15.36 | 11.79 | 14 | 12.5803 | 15.75 | 15.5453 | 18.4 | 15.6 | 14.59 | 15.6 |  |  |  | 1
| 1 | 2 | 16.9 | 15.23 | 14.0155 | 14.567 | 12.4 | 17.3 | 14.9 | 15.81 | 18 | 15 | 14.9 | 16.15 | 15.3 | 15.55 | 10 | 16.1 | 15.99 | 13.74 | 14 | 12.192 | 14.15 | 15.7768 | 18.1 | 15.3 | 14.85 | 15.1 |  |  |  | 32
| 1 | 3 | 16.8 | 15.07 | 13.8256 | 13.5571 | 12.7 | 15.8 | 15.3 | 15.25 | 18 | 15 | 14.75 | 15.45 | 14.8 | 15.2 | 10 | 16.5 | 15.51 | 10.87 | 14 | 13.017 | 15.4 | 15.8536 | 18.1 | 15.3 | 14.83 | 14.8 |  |  |  | 16
| 1 | 4 | 16.7 | 14.6 | 13.8545 | 13.8366 | 12.6 | 15.7 | 15.2 | 15.13 | 18 | 15 | 14.85 | 16.65 | 14.6 | 14.6 | 11 | 15.1 | 15.1 | 12.77 | 14 | 12.8684 | 15.2 | 16.7303 | 18.2 | 15.2 | 14.53 | 15 |  |  |  | 15.334
| 1 | 5 | 17.4 | 15.08 | 13.9817 | 13.6643 | 12.7 | 15.6 | 15.4 | 16.03 | 17 | 15 | 14.9 | 17.1 | 15.2 | 15.25 | 11 | 16.6 | 15.77 | 13.81 | 13 | 13.4509 | 15.7 | 15.7407 | 18.9 | 15.5 | 15.03 | 15.7 |  |  |  | 30
| 1 | 6 | 17.1 | 14.71 | 14.2239 | 14.3703 | 13.4 | 15 | 15.9 | 16.07 | 18 | 15 | 14.75 | 16.65 | 15.15 | 14.75 | 11 | 15.9 | 15.69 | 11.57 | 14 | 13.2575 | 15.2 | 15.7026 | 18.4 | 16.4 | 14.54 | 16.7
| Mean |  | 16.9333 | 14.9383 | 13.9207 | 13.9579 | 12.7833 | 16.2333 | 15.3167 | 15.5717 | 17.8333 | 15 | 14.8583 | 16.5167 | 15.0333 | 15.05 | 10.5 | 15.9167 | 15.57 | 12.425 | 13.8333 | 12.8943 | 15.2333 | 15.8916 | 18.35 | 15.55 | 14.7283 | 15.4833
| Median |  | 16.85 | 15.005 | 13.9181 | 13.7942 | 12.7 | 15.75 | 15.25 | 15.53 | 18 | 15 | 14.875 | 16.65 | 15.15 | 15.075 | 10.5 | 16 | 15.6 | 12.28 | 14 | 12.9427 | 15.3 | 15.7588 | 18.3 | 15.4 | 14.71 | 15.35
| Std Dev. |  | 0.27325 | 0.24045 | 0.20338 | 0.41111 | 0.34303 | 1.1535 | 0.33116 | 0.44723 | 0.40825 | 0 | 0.09704 | 0.63061 | 0.27142 | 0.35071 | 0.54772 | 0.61455 | 0.31604 | 1.2099 | 0.40825 | 0.45807 | 0.58109 | 0.42343 | 0.30166 | 0.44159 | 0.20498 | 0.69113
| Rel.Std.Dev. |  | 1.614% | 1.61% | 1.461% | 2.945% | 2.683% | 7.106% | 2.162% | 2.872% | 2.289% | 0% | 0.6531% | 3.818% | 1.805% | 2.33% | 5.216% | 3.861% | 2.03% | 9.737% | 2.951% | 3.553% | 3.815% | 2.664% | 1.644% | 2.84% | 1.392% | 4.464%
| PDM3 |  | 10.43% | -2.58% | -9.217% | -8.975% | -16.63% | 5.865% | -0.1132% | 1.55% | 16.3% | -2.178% | -3.102% | 7.713% | -1.961% | -1.852% | -31.52% | 3.8% | 1.539% | -18.97% | -9.787% | -15.91% | -0.6566% | 3.636% | 19.67% | 1.409% | -3.95% | 0.9738%
| Z-Score (Absolute) |  | 1.92 | 0.39 | 1.57 | 1.53 | 2.89 | 1.11 | 0.05 | 0.34 | NA | NA | 0.48 | 1.44 | 0.28 | 0.26 | NA | 0.74 | 0.34 | 3.3 | NA | 2.76 | 0.05 | 0.71 | 3.57 | 0.32 | 0.63 | 0.24
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 241b (Certified Value 3.06 ppm) | [30 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 3.14 | 2.69 | 3.242 | 2.9 | 3.3 | 3.21 | 3.2 | 0.55058 | [21 empty] | 1
| 1 | 2 | 3.21 | 2.74 | 3.2846 | 2.8 | 3.2 | 3.13 | 3.2 | 0.57881 | [21 empty] | 33
| 1 | 3 | 3.17 | 2.72 | 3.3956 | 3 | 2.9 | 3.13 | 3.2 | 0.56368 | [21 empty] | 16
| 1 | 4 | 3.23 | 2.66 | 3.2926 | 2.8 | 2.9 | 3.08 | 3.1 | 0.59777 | [21 empty] | 3.058
| 1 | 5 | 3.2 | 2.79 | 3.3058 | 2.9 | 2.9 | 3.06 | 3.1 | 0.57738 | [21 empty] | 31
| 1 | 6 | 3.15 | 2.69 | 3.3264 | 3 | 2.8 | 3.19 | 3.2 | 0.57308
| Mean |  | 3.1833 | 2.715 | 3.3078 | 2.9 | 3 | 3.1333 | 3.1667 | 0.57355
| Median |  | 3.185 | 2.705 | 3.2992 | 2.9 | 2.9 | 3.13 | 3.2 | 0.57523
| Std Dev. |  | 0.03559 | 0.045935 | 0.051295 | 0.089443 | 0.2 | 0.058878 | 0.05164 | 0.015838
| Rel.Std.Dev. |  | 1.118% | 1.692% | 1.551% | 3.084% | 6.667% | 1.879% | 1.631% | 2.761%
| PDM3 |  | 4.098% | -11.22% | 8.169% | -5.168% | -1.897% | 2.463% | 3.553% | -81.24%
| Z-Score (Absolute) |  | 0.56 | 1.67 | 1.15 | 0.79 | 0.32 | 0.32 | 0.48 | 11.87
| Analytical results for Ge in OREAS 241b (Certified Value 0.1 ppm) | [30 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
2
| 1 | 1 | < 0.1 | 1.6 | 0.15141 | < 0.2 | 0.2 | 1.6 | 1.1 | 2.9 | 0.08 | 0.1 | 0.05 | 0.07 | < 10 | 0.8 | < 5 | 0.08 | 0.1233 | 0.13 | 0.2 | 0.05 | 0.12 | [8 empty] | 1
| 1 | 2 | < 0.1 | 1.8 | 0.15175 | < 0.2 | 0.1 | 1.9 | 1.2 | 3.1 | 0.08 | 0.09 | 0.07 | 0.07 | < 10 | 0.8 | < 5 | 0.26 | 0.1701 | 0.1 | 0.2 | < 0.05 | 0.12 | [8 empty] | 34
| 1 | 3 | < 0.1 | 1.6 | 0.1474 | < 0.2 | 0.2 | 1.7 | 1.2 | 2.9 | 0.1 | 0.09 | 0.08 | 0.07 | < 10 | 0.7 | < 5 | 0.09 | 0.1677 | 0.11 | 0.1 | < 0.05 | 0.18 | [8 empty] | 16
| 1 | 4 | < 0.1 | 1.7 | 0.15562 | < 0.2 | 0.2 | 1.3 | 1.2 | 3 | 0.09 | 0.1 | 0.08 | 0.07 | < 10 | 0.6 | < 5 | 0.07 | 0.1586 | 0.14 | 0.3 | < 0.05 | 0.12 | [8 empty] | 0.10253
| 1 | 5 | 0.2 | 1.7 | 0.15435 | < 0.2 | 0.2 | 1.9 | 1.2 | 3 | 0.08 | 0.1 | 0.07 | 0.07 | < 10 | 0.7 | < 5 | 0.06 | 0.1559 | 0.13 | 0.2 | < 0.05 | 0.12 | [8 empty] | 32
| 1 | 6 | < 0.1 | 1.6 | 0.14677 | < 0.2 | 0.3 | 1.4 | 1.1 | 3.1 | 0.07 | 0.1 | 0.09 | 0.08 | < 10 | 0.7 | < 5 | 0.2 | 0.16 | 0.09 | 0.1 | < 0.05 | 0.12
| Mean |  | 0.2 | 1.6667 | 0.15121 |  | 0.2 | 1.6333 | 1.1667 | 3 | 0.083333 | 0.096667 | 0.073333 | 0.071667 |  | 0.71667 |  | 0.12667 | 0.15593 | 0.11667 | 0.18333 | 0.05 | 0.13
| Median |  | 0.2 | 1.65 | 0.15158 |  | 0.2 | 1.65 | 1.2 | 3 | 0.08 | 0.1 | 0.075 | 0.07 |  | 0.7 |  | 0.085 | 0.1593 | 0.12 | 0.2 | 0.05 | 0.12
| Std Dev. |  |  | 0.08165 | 0.0035735 |  | 0.063246 | 0.25033 | 0.05164 | 0.089443 | 0.010328 | 0.005164 | 0.013663 | 0.0040825 |  | 0.075277 |  | 0.082865 | 0.016898 | 0.019664 | 0.075277 |  | 0.024495
| Rel.Std.Dev. |  |  | 4.899% | 2.363% |  | 31.62% | 15.33% | 4.426% | 2.981% | 12.39% | 5.342% | 18.63% | 5.696% |  | 10.5% |  | 65.42% | 10.84% | 16.85% | 41.06% |  | 18.84%
| PDM3 |  | 95.06% | 1525% | 47.48% |  | 95.06% | 1493% | 1038% | 2826% | -18.73% | -5.722% | -28.48% | -30.1% |  | 599% |  | 23.54% | 52.08% | 13.78% | 78.8% | -51.24% | 26.79%
| Z-Score (Absolute) |  | NA | NA | 0.67 | 0.33 | NA | NA | NA | NA | 0.65 | 0.39 | 0.85 | 0.88 | 95.31 | NA | 46.52 | 0.2 | 0.77 | 0 | NA | 1.71 | 0.26
Results from laboratories 1.01, 1.02, 1.07, 1.10, 1.11, 1.13, 1.22 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 241b (Certified Value 1.75 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 1.84 | 1.71 | 0.79983 | 1.3 | 2.1 | 1.7 | 1.63 | 1.9 | 1.7 | 1.8 | 1.7 | 1.6 | 1.7 | 1.5 | 1.93 | 2 | 2.0608 | 1.7 | 1.5707 | 1.8 | 2.33 | 2 | 1.72 | [6 empty] | 1
| 1 | 2 | 1.86 | 1.75 | 0.77076 | 1.6 | 2 | 1.8 | 1.77 | 1.9 | 1.6 | 1.8 | 1.7 | 1.6 | 1.8 | 1.5 | 1.89 | 2 | 2.1177 | 1.7 | 1.5539 | 1.8 | 2.52 | 1.9 | 1.66 | [6 empty] | 18
| 1 | 3 | 1.88 | 1.66 | 0.75455 | 1 | 1.9 | 1.8 | 1.89 | 1.8 | 1.7 | 1.7 | 1.6 | 1.7 | 1.8 | 1.5 | 1.83 | 2 | 2.1402 | 1.7 | 1.5673 | 1.8 | 2.28 | 2.1 | 1.63 | [6 empty] | 16
| 1 | 4 | 1.78 | 1.67 | 0.78081 | 1.1 | 1.9 | 1.8 | 1.7 | 1.9 | 1.6 | 1.8 | 1.7 | 1.6 | 1.7 | 1.5 | 1.81 | 2 | 2.0988 | 1.7 | 1.5686 | 1.8 | 2.33 | 1.8 | 1.67 | [6 empty] | 1.7504
| 1 | 5 | 1.88 | 1.66 | 0.75681 | 1.4 | 1.9 | 1.8 | 1.67 | 1.9 | 1.7 | 1.9 | 1.7 | 1.8 | 1.8 | 1.5 | 1.8 | 2 | 2.1199 | 1.8 | 1.5205 | 2.8 | 2.33 | 1.8 | 1.73 | [6 empty] | 33
| 1 | 6 | 1.94 | 1.62 | 0.78751 | 1.3 | 1.8 | 1.8 | 1.81 | 2 | 1.6 | 1.8 | 1.7 | 1.7 | 1.7 | 1.5 | 1.8 | 2 | 2.1037 | 1.7 | 1.5638 | 2 | 2.47 | 1.8 | 1.77
| Mean |  | 1.8633 | 1.6783 | 0.77505 | 1.2833 | 1.9333 | 1.7833 | 1.745 | 1.9 | 1.65 | 1.8 | 1.6833 | 1.6667 | 1.75 | 1.5 | 1.8433 | 2 | 2.1069 | 1.7167 | 1.5575 | 2 | 2.3767 | 1.9 | 1.6967
| Median |  | 1.87 | 1.665 | 0.77579 | 1.3 | 1.9 | 1.8 | 1.735 | 1.9 | 1.65 | 1.8 | 1.7 | 1.65 | 1.75 | 1.5 | 1.82 | 2 | 2.1107 | 1.7 | 1.5655 | 1.8 | 2.33 | 1.85 | 1.695
| Std Dev. |  | 0.052789 | 0.04535 | 0.017741 | 0.2137 | 0.10328 | 0.040825 | 0.096695 | 0.063246 | 0.054772 | 0.063246 | 0.040825 | 0.08165 | 0.054772 | 0 | 0.054283 | 0 | 0.026808 | 0.040825 | 0.019053 | 0.4 | 0.095009 | 0.12649 | 0.052026
| Rel.Std.Dev. |  | 2.833% | 2.702% | 2.289% | 16.65% | 5.342% | 2.289% | 5.541% | 3.329% | 3.32% | 3.514% | 2.425% | 4.899% | 3.13% | 0% | 2.945% | 0% | 1.272% | 2.378% | 1.223% | 20% | 3.998% | 6.657% | 3.066%
| PDM3 |  | 6.451% | -4.118% | -55.72% | -26.68% | 10.45% | 1.881% | -0.3093% | 8.546% | -5.737% | 2.833% | -3.832% | -4.784% | -0.0237% | -14.31% | 5.308% | 14.26% | 20.36% | -1.928% | -11.02% | 14.26% | 35.78% | 8.546% | -3.071%
| Z-Score (Absolute) |  | 0.73 | 0.44 | 6.15 | 2.93 | 1.17 | 0.23 | 0.02 | 0.96 | 0.62 | 0.33 | 0.41 | 0.51 | 0.02 | 1.56 | 0.61 | NA | 2.27 | 0.19 | 1.2 | 1.6 | 3.98 | 0.96 | 0.32
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 241b (Indicative Value 0.077 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.05 | 1.25 | 1.30 | [26 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | [26 empty] | 3
3
| 1 | 1 | < 1 | < 0.1 | 0.0179 | [26 empty] | 1
| 1 | 2 | < 1 | < 0.1 | 0.0356 | [26 empty] | 5
| 1 | 3 | < 1 | < 0.1 | 0.0187 | [26 empty] | 16
| 1 | 4 | < 1 | 0.4739 | 0.023 | [26 empty] | 0.076987
| 1 | 5 | < 1 | 0.4717 | < 0.01 | [26 empty] | 9
| 1 | 6 | < 1 | < 0.1 | 0.0176
| Mean |  |  | 0.4728 | 0.02256
| Median |  |  | 0.4728 | 0.0187
| Std Dev. |  |  | 0.0015556 | 0.0076068
| Rel.Std.Dev. |  |  | 0.329% | 33.72%
| PDM3 |  |  | 514.1% | -70.7%
| Z-Score (Absolute) |  | 1.22 | 0 | 0.67
| Analytical results for Ho in OREAS 241b (Certified Value 0.76 ppm) | [30 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 0.77 | 0.67 | 0.65441 | 0.8 | 0.83 | 0.76 | 0.9 | 0.79547 | [21 empty] | 1
| 1 | 2 | 0.78 | 0.69 | 0.69479 | 0.7 | 0.81 | 0.77 | 0.75 | 0.80701 | [21 empty] | 19
| 1 | 3 | 0.76 | 0.67 | 0.68197 | 0.7 | 0.76 | 0.77 | 0.8 | 0.8255 | [21 empty] | 16
| 1 | 4 | 0.79 | 0.67 | 0.69712 | 0.8 | 0.75 | 0.78 | 0.8 | 0.82593 | [21 empty] | 0.75788
| 1 | 5 | 0.8 | 0.69 | 0.67992 | 0.7 | 0.74 | 0.76 | 0.85 | 0.81181 | [21 empty] | 34
| 1 | 6 | 0.75 | 0.66 | 0.69158 | 0.7 | 0.71 | 0.77 | 0.85 | 0.80556
| Mean |  | 0.775 | 0.675 | 0.6833 | 0.73333 | 0.76667 | 0.76833 | 0.825 | 0.81188
| Median |  | 0.775 | 0.67 | 0.68677 | 0.7 | 0.755 | 0.77 | 0.825 | 0.80941
| Std Dev. |  | 0.018708 | 0.012247 | 0.015743 | 0.05164 | 0.045019 | 0.0075277 | 0.05244 | 0.011965
| Rel.Std.Dev. |  | 2.414% | 1.814% | 2.304% | 7.042% | 5.872% | 0.9797% | 6.356% | 1.474%
| PDM3 |  | 2.259% | -10.94% | -9.841% | -3.239% | 1.159% | 1.379% | 8.856% | 7.125%
| Z-Score (Absolute) |  | 0.1 | 1.45 | 1.32 | NA | 0.03 | 0 | 0.88 | 0.67
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 241b (Certified Value 0.078 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.08 | 0.08 | 0.057633 | < 0.1 | 0.07 | 0.07 | 0.08 | 0.08 | 0.078 | 0.081 | 0.081 | < 10 | 0.09 | 0.07 | 0.075 | 25.4 | < 0.1 | 0.06932 | 0.048 | 0.0773 | 0.077 | 0.08 | 0.076 | 0.078 | [5 empty] | 1
| 1 | 2 | 0.08 | 0.08 | 0.066331 | < 0.1 | 0.07 | 0.07 | 0.08 | 0.085 | 0.079 | 0.083 | 0.08 | < 10 | 0.09 | 0.07 | 0.077 | 25.69 | 0.1 | 0.06708 | 0.07 | 0.0807 | 0.078 | 0.08 | 0.075 | 0.08 | [5 empty] | 20
| 1 | 3 | 0.09 | 0.07 | 0.058627 | < 0.1 | 0.07 | 0.08 | 0.09 | 0.079 | 0.075 | 0.076 | 0.08 | < 10 | 0.09 | 0.07 | 0.075 | 26.06 | 0.1 | 0.07133 | 0.057 | 0.077 | 0.07 | 0.08 | 0.077 | 0.078 | [5 empty] | 16
| 1 | 4 | 0.09 | 0.08 | 0.063598 | < 0.1 | 0.08 | 0.08 | 0.09 | 0.081 | 0.077 | 0.08 | 0.078 | < 10 | 0.09 | 0.07 | 0.073 | 23.39 | 0.1 | 0.06863 | 0.053 | 0.0826 | 0.073 | 0.07 | 0.072 | 0.077 | [5 empty] | 0.078107
| 1 | 5 | 0.09 | 0.07 | 0.06459 | < 0.1 | 0.06 | 0.08 | 0.08 | 0.082 | 0.078 | 0.085 | 0.082 | < 10 | 0.09 | 0.07 | 0.073 | 23.38 | < 0.1 | 0.07305 | 0.057 | 0.0742 | 0.078 | 0.08 | 0.072 | 0.082 | [5 empty] | 35
| 1 | 6 | 0.08 | 0.09 | 0.063411 | < 0.1 | 0.08 | 0.08 | 0.09 | 0.078 | 0.081 | 0.078 | 0.081 | < 10 | 0.08 | 0.07 | 0.076 | 24.04 | 0.1 | 0.07148 | 0.071 | 0.0739 | 0.085 | 0.08 | 0.071 | 0.083
| Mean |  | 0.085 | 0.078333 | 0.062365 |  | 0.071667 | 0.076667 | 0.085 | 0.080833 | 0.078 | 0.0805 | 0.080333 |  | 0.088333 | 0.07 | 0.074833 | 24.66 | 0.1 | 0.070148 | 0.059333 | 0.077617 | 0.076833 | 0.078333 | 0.073833 | 0.079667
| Median |  | 0.085 | 0.08 | 0.063505 |  | 0.07 | 0.08 | 0.085 | 0.0805 | 0.078 | 0.0805 | 0.0805 |  | 0.09 | 0.07 | 0.075 | 24.72 | 0.1 | 0.070325 | 0.057 | 0.07715 | 0.0775 | 0.08 | 0.0735 | 0.079
| Std Dev. |  | 0.0054772 | 0.0075277 | 0.0034541 |  | 0.0075277 | 0.005164 | 0.0054772 | 0.0024833 | 0.002 | 0.0032711 | 0.0013663 |  | 0.0040825 | 0 | 0.0016021 | 1.2003 | 0 | 0.0021907 | 0.0092664 | 0.003473 | 0.0051153 | 0.0040825 | 0.0024833 | 0.0024221
| Rel.Std.Dev. |  | 6.444% | 9.61% | 5.538% |  | 10.5% | 6.736% | 6.444% | 3.072% | 2.564% | 4.063% | 1.701% |  | 4.622% | 0% | 2.141% | 4.868% | 0% | 3.123% | 15.62% | 4.475% | 6.658% | 5.212% | 3.363% | 3.04%
| PDM3 |  | 8.825% | 0.2894% | -20.15% |  | -8.246% | -1.844% | 8.825% | 3.49% | -0.1374% | 3.063% | 2.85% |  | 13.09% | -10.38% | -4.192% | 31472% | 28.03% | -10.19% | -24.04% | -0.6282% | -1.631% | 0.2894% | -5.472% | 1.996%
| Z-Score (Absolute) |  | 0.97 | 0.02 | 2.24 | 4 | 0.92 | 0.21 | 0.97 | 0.38 | 0.02 | 0.33 | 0.31 | 698.7 | 1.44 | 1.16 | 0.47 | 3489.63 | 0.73 | 1.14 | 2.67 | 0.08 | 0.19 | 0.02 | 0.62 | 0.21
| Analytical results for K in OREAS 241b (Certified Value 0.629 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.64 | 0.6076 | 0.61472 | 0.61949 | 0.61 | 0.6691 | 0.596 | 0.5863 | 0.62 | 0.67 | 0.62 | 0.68 | 0.6 | 0.6 | 0.64 | 0.637 | 0.57 | 0.65 | 0.79 | 0.63952 | 0.605 | 0.61 | 0.6802 | 0.61 | 0.67 | 0.65 | 0.65 |  |  | 1
| 1 | 2 | 0.64 | 0.6198 | 0.61774 | 0.63727 | 0.59 | 0.6791 | 0.607 | 0.6061 | 0.63 | 0.66 | 0.62 | 0.65 | 0.6 | 0.6 | 0.6 | 0.627 | 0.57 | 0.66 | 0.903 | 0.63128 | 0.6 | 0.58 | 0.6644 | 0.61 | 0.67 | 0.65 | 0.64 |  |  | #N/A
| 1 | 3 | 0.65 | 0.6253 | 0.61695 | 0.61549 | 0.59 | 0.6782 | 0.617 | 0.5941 | 0.62 | 0.66 | 0.64 | 0.64 | 0.59 | 0.62 | 0.62 | 0.631 | 0.58 | 0.65 | 1.048 | 0.63517 | 0.605 | 0.62 | 0.6911 | 0.63 | 0.68 | 0.66 | 0.63 |  |  | 16
| 1 | 4 | 0.66 | 0.6128 | 0.61772 | 0.62306 | 0.59 | 0.645 | 0.601 | 0.5974 | 0.62 | 0.65 | 0.65 | 0.66 | 0.59 | 0.6 | 0.65 | 0.617 | 0.57 | 0.66 | 1.341 | 0.63354 | 0.59 | 0.61 | 0.67 | 0.63 | 0.66 | 0.65 | 0.65 |  |  | 0.62862
| 1 | 5 | 0.66 | 0.6244 | 0.61313 | 0.61806 | 0.58 | 0.6658 | 0.612 | 0.6098 | 0.63 | 0.65 | 0.62 | 0.67 | 0.65 | 0.61 | 0.64 | 0.62 | 0.58 | 0.64 | 1.541 | 0.62453 | 0.605 | 0.63 | 0.6928 | 0.63 | 0.68 | 0.65 | 0.63 |  |  | 36
| 1 | 6 | 0.64 | 0.6082 | 0.61734 | 0.62811 | 0.63 | 0.6733 | 0.619 | 0.6088 | 0.61 | 0.66 | 0.61 | 0.65 | 0.6 | 0.61 | 0.66 | 0.615 | 0.57 | 0.67 | 1.185 | 0.62638 | 0.6 | 0.6 | 0.666 | 0.63 | 0.71 | 0.65 | 0.63
| Mean |  | 0.64833 | 0.61635 | 0.61627 | 0.62358 | 0.59833 | 0.66842 | 0.60867 | 0.60042 | 0.62167 | 0.65833 | 0.62667 | 0.65833 | 0.605 | 0.60667 | 0.635 | 0.6245 | 0.57333 | 0.655 | 1.1347 | 0.63174 | 0.60083 | 0.60833 | 0.67742 | 0.62333 | 0.67833 | 0.65167 | 0.63833
| Median |  | 0.645 | 0.6163 | 0.61714 | 0.62128 | 0.59 | 0.6712 | 0.6095 | 0.60175 | 0.62 | 0.66 | 0.62 | 0.655 | 0.6 | 0.605 | 0.64 | 0.6235 | 0.57 | 0.655 | 1.1165 | 0.63241 | 0.6025 | 0.61 | 0.6751 | 0.63 | 0.675 | 0.65 | 0.635
| Std Dev. |  | 0.0098319 | 0.0079044 | 0.0019037 | 0.0080102 | 0.018348 | 0.012565 | 0.009048 | 0.0093692 | 0.0075277 | 0.0075277 | 0.015055 | 0.01472 | 0.022583 | 0.008165 | 0.021679 | 0.0086197 | 0.005164 | 0.010488 | 0.27935 | 0.0055924 | 0.0058452 | 0.017224 | 0.01254 | 0.010328 | 0.017224 | 0.0040825 | 0.0098319
| Rel.Std.Dev. |  | 1.516% | 1.282% | 0.3089% | 1.285% | 3.067% | 1.88% | 1.487% | 1.56% | 1.211% | 1.143% | 2.402% | 2.236% | 3.733% | 1.346% | 3.414% | 1.38% | 0.9007% | 1.601% | 24.62% | 0.8852% | 0.9729% | 2.831% | 1.851% | 1.657% | 2.539% | 0.6265% | 1.54%
| PDM3 |  | 3.136% | -1.952% | -1.966% | -0.8024% | -4.818% | 6.33% | -3.175% | -4.487% | -1.107% | 4.726% | -0.3111% | 4.726% | -3.758% | -3.493% | 1.015% | -0.6558% | -8.795% | 4.196% | 80.5% | 0.4954% | -4.421% | -3.228% | 7.762% | -0.8414% | 7.908% | 3.666% | 1.545%
| Z-Score (Absolute) |  | 0.82 | 0.28 | 0.28 | 0.03 | 0.9 | 1.52 | 0.55 | 0.83 | 0.1 | 1.17 | 0.07 | 1.17 | 0.67 | 0.62 | 0.36 | 0 | 1.77 | 1.05 | 17.64 | 0.25 | 0.82 | 0.56 | 1.83 | 0.04 | 1.86 | 0.94 | 0.48
| Analytical results for La in OREAS 241b (Certified Value 6.34 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 8.1 | 6.36 | 5.64 | 6.41 | 6.3962 | 6.1 | 7 | 6.06 | 6 | 10 | 6.2 | 6.9 | 6.4 | 6.4 | 7 | 5.9 | 6 | 7 | 6.1719 | 6.6 | 5.8151 | 7.8 | 6.3 | 6.7 | 7.2 |  |  |  |  | 1
| 1 | 2 | 8 | 6.31 | 5.83 | 6.43 | 6.6456 | 6 | 7 | 6.33 | 6 | 9 | 6.1 | 7 | 6.6 | 6.3 | 6 | 5.9 | 6.3 | 6.5 | 6.3949 | 6.2 | 5.8437 | 7.8 | 6.2 | 6.6 | 7 |  |  |  |  | 21
| 1 | 3 | 8.3 | 6.32 | 5.87 | 6.41 | 6.2633 | 6 | 7 | 6.51 | 6 | 9 | 6.1 | 6.7 | 6.2 | 6 | 6 | 5.7 | 6 | 7 | 6.4306 | 6.5 | 5.8355 | 7.7 | 6.3 | 6.7 | 6.9 |  |  |  |  | 16
| 1 | 4 | 6 | 6.23 | 5.71 | 6.45 | 6.3661 | 6 | 6 | 6.14 | 6 | 10 | 6.1 | 7.1 | 6.2 | 6.1 | 7 | 5.8 | 6.1 | 7 | 6.4084 | 6.4 | 5.8323 | 7.8 | 6.3 | 6.2 | 7.1 |  |  |  |  | 6.3353
| 1 | 5 | 8.4 | 6.36 | 5.84 | 6.57 | 6.3433 | 5.9 | 7 | 6.18 | 6 | 9 | 6.1 | 7.3 | 6.4 | 6.1 | 7 | 6 | 6 | 6.5 | 6.3402 | 6.5 | 5.6791 | 7.8 | 6.6 | 6.3 | 7.4 |  |  |  |  | 37
| 1 | 6 | 7.9 | 6.21 | 5.78 | 6.52 | 6.6334 | 6.3 | 7 | 6.3 | 6 | 10 | 6.5 | 7.3 | 6.3 | 6 | 7 | 5.7 | 6.4 | 7 | 6.3666 | 6.5 | 5.8325 | 7.7 | 6.7 | 6.2 | 7.8
| Mean |  | 7.7833 | 6.2983 | 5.7783 | 6.465 | 6.4413 | 6.05 | 6.8333 | 6.2533 | 6 | 9.5 | 6.1833 | 7.05 | 6.35 | 6.15 | 6.6667 | 5.8333 | 6.1333 | 6.8333 | 6.3521 | 6.45 | 5.8064 | 7.7667 | 6.4 | 6.45 | 7.2333
| Median |  | 8.05 | 6.315 | 5.805 | 6.44 | 6.3812 | 6 | 7 | 6.24 | 6 | 9.5 | 6.1 | 7.05 | 6.35 | 6.1 | 7 | 5.85 | 6.05 | 7 | 6.3807 | 6.5 | 5.8324 | 7.8 | 6.3 | 6.45 | 7.15
| Std Dev. |  | 0.89312 | 0.064317 | 0.087958 | 0.065651 | 0.15974 | 0.13784 | 0.40825 | 0.16096 | 0 | 0.54772 | 0.16021 | 0.23452 | 0.15166 | 0.16432 | 0.5164 | 0.12111 | 0.17512 | 0.2582 | 0.093795 | 0.13784 | 0.063041 | 0.05164 | 0.2 | 0.2429 | 0.3266
| Rel.Std.Dev. |  | 11.47% | 1.021% | 1.522% | 1.015% | 2.48% | 2.278% | 5.974% | 2.574% | 0% | 5.766% | 2.591% | 3.327% | 2.388% | 2.672% | 7.746% | 2.076% | 2.855% | 3.779% | 1.477% | 2.137% | 1.086% | 0.6649% | 3.125% | 3.766% | 4.515%
| PDM3 |  | 22.86% | -0.5836% | -8.792% | 2.047% | 1.673% | -4.503% | 7.861% | -1.294% | -5.293% | 49.95% | -2.399% | 11.28% | 0.232% | -2.925% | 5.23% | -7.923% | -3.188% | 7.861% | 0.2645% | 1.81% | -8.349% | 22.59% | 1.021% | 1.81% | 14.17%
| Z-Score (Absolute) |  | 4.78 | 0.18 | 1.91 | 0.38 | 0.3 | 1.01 | NA | 0.33 | NA | NA | 0.56 | 2.33 | 0.01 | 0.67 | NA | 1.73 | 0.73 | 1.61 | 0 | 0.33 | 1.82 | 4.72 | 0.16 | 0.33 | 2.94
Results from laboratories 1.08, 1.12, 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 241b (Certified Value 12.8 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | < 1 | 12.4 | 12.16 | 5.9513 | 13.0197 | 12 | 6 | 13.1 | 11.8 | 12 | 12.7 | 12.4 | 13.7 | 13.2 | 13 | < 5 | 12.8 | 10.9 | 12 | 13 | 10.4895 | 12.4 | 10.8654 | 12.6 | 15 | 10.1 | 12.1 |  |  | 1
| 1 | 2 | < 1 | 12.6 | 11.87 | 6.1034 | 13.5216 | 12 | 6 | 13.2 | 12.1 | 11 | 12.6 | 12.6 | 13.2 | 13 | 12.7 | < 5 | 12.5 | 10.8 | 12 | 13 | 9.9237 | 12.2 | 10.7328 | 12.4 | 15 | 10.2 | 12.2 |  |  | #N/A
| 1 | 3 | < 1 | 12.2 | 14.03 | 6.0743 | 12.8902 | 12 | 6 | 13.1 | 12.3 | 12 | 13.2 | 13.2 | 13.1 | 12.8 | 13.1 | < 5 | 12.4 | 10.9 | 12 | 14 | 10.6669 | 12.6 | 11.0058 | 12.8 | 15 | 10.2 | 11.7 |  |  | 16
| 1 | 4 | < 1 | 12.1 | 13.46 | 5.8629 | 13.08 | 12 | 6 | 13.4 | 12.2 | 12 | 14.2 | 13.7 | 13.5 | 12.8 | 12.7 | < 5 | 12.2 | 11 | 12 | 12 | 10.4516 | 13 | 10.6548 | 13 | 15 | 9.9 | 12.1 |  |  | 12.8257
| 1 | 5 | < 1 | 12.6 | 13.6 | 5.9601 | 12.9428 | 12 | 6 | 13.6 | 12 | 12 | 14.6 | 12.2 | 13.8 | 14.2 | 12.8 | < 5 | 12.4 | 11.2 | 12 | 13 | 10.8179 | 13.2 | 11.1072 | 12.8 | 15 | 10.2 | 12.9 |  |  | 38
| 1 | 6 | < 1 | 12 | 13.41 | 5.9899 | 13.2576 | 12 | 6 | 13.2 | 12.2 | 12 | 14.1 | 12.6 | 13.4 | 13 | 13 | < 5 | 12.8 | 10.9 | 12 | 13 | 10.7113 | 12.8 | 10.6392 | 13 | 16 | 9.9 | 12.5
| Mean |  |  | 12.3167 | 13.0883 | 5.9903 | 13.1187 | 12 | 6 | 13.2667 | 12.1 | 11.8333 | 13.5667 | 12.7833 | 13.45 | 13.1667 | 12.8833 |  | 12.5167 | 10.95 | 12 | 13 | 10.5102 | 12.7 | 10.8342 | 12.7667 | 15.1667 | 10.0833 | 12.25
| Median |  |  | 12.3 | 13.435 | 5.975 | 13.0499 | 12 | 6 | 13.2 | 12.15 | 12 | 13.65 | 12.6 | 13.45 | 13 | 12.9 |  | 12.45 | 10.9 | 12 | 13 | 10.5782 | 12.7 | 10.7991 | 12.8 | 15 | 10.15 | 12.15
| Std Dev. |  |  | 0.25626 | 0.86444 | 0.087775 | 0.23496 | 0 | 0 | 0.19664 | 0.17889 | 0.40825 | 0.84538 | 0.56006 | 0.27386 | 0.52789 | 0.17224 |  | 0.24014 | 0.13784 | 0 | 0.63246 | 0.31861 | 0.37417 | 0.19258 | 0.23381 | 0.40825 | 0.1472 | 0.40866
| Rel.Std.Dev. |  |  | 2.081% | 6.605% | 1.465% | 1.791% | 0% | 0% | 1.482% | 1.478% | 3.45% | 6.231% | 4.381% | 2.036% | 4.009% | 1.337% |  | 1.919% | 1.259% | 0% | 4.865% | 3.031% | 2.946% | 1.778% | 1.831% | 2.692% | 1.46% | 3.336%
| PDM3 |  |  | -3.969% | 2.047% | -53.29% | 2.284% | -6.438% | -53.22% | 3.438% | -5.658% | -7.738% | 5.777% | -0.3306% | 4.867% | 2.658% | 0.4491% |  | -2.41% | -14.62% | -6.438% | 1.359% | -18.05% | -0.9804% | -15.53% | -0.4606% | 18.25% | -21.38% | -4.489%
| Z-Score (Absolute) |  | 12.47 | 0.21 | 0.59 | 6.77 | 0.62 | NA | NA | 0.78 | 0.43 | NA | 1.09 | 0.28 | 0.97 | 0.67 | 0.38 | 10.39 | 0 | 1.63 | NA | NA | 2.08 | 0.19 | 1.75 | 0.26 | NA | 2.52 | 0.28
Results from laboratories 1.07, 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 241b (Certified Value 0.32 ppm) | [30 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [20 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | 1.32 | [20 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [20 empty] | 2
3
| 1 | 1 | 0.34 | 0.32 | 0.28 | 4.0228 | 0.3 | 0.35 | 0.5 | 0.33111 | 0.29 | [20 empty] | 1
| 1 | 2 | 0.35 | 0.32 | 0.27 | 4.0673 | 0.3 | 0.33 | 0.5 | 0.34822 | 0.29 | [20 empty] | 22
| 1 | 3 | 0.34 | 0.33 | 0.27 | 4.1155 | 0.3 | 0.33 | 0.5 | 0.35042 | 0.29 | [20 empty] | 16
| 1 | 4 | 0.35 | 0.32 | 0.27 | 4.0649 | 0.3 | 0.32 | 0.5 | 0.35984 | 0.28 | [20 empty] | 0.31966
| 1 | 5 | 0.36 | 0.32 | 0.28 | 4.0915 | 0.3 | 0.33 | 0.5 | 0.34353 | 0.31 | [20 empty] | 39
| 1 | 6 | 0.35 | 0.32 | 0.27 | 4.0638 | 0.3 | 0.33 | 0.5 | 0.3445 | 0.32
| Mean |  | 0.34833 | 0.32167 | 0.27333 | 4.071 | 0.3 | 0.33167 | 0.5 | 0.34627 | 0.29667
| Median |  | 0.35 | 0.32 | 0.27 | 4.0661 | 0.3 | 0.33 | 0.5 | 0.34636 | 0.29
| Std Dev. |  | 0.0075277 | 0.0040825 | 0.005164 | 0.031083 | 0 | 0.0098319 | 0 | 0.0094407 | 0.015055
| Rel.Std.Dev. |  | 2.161% | 1.269% | 1.889% | 0.7635% | 0% | 2.964% | 0% | 2.726% | 5.075%
| PDM3 |  | 8.971% | 0.629% | -14.49% | 1174% | -6.149% | 3.757% | 56.42% | 8.326% | -7.192%
| Z-Score (Absolute) |  | 0.67 | 0.4 | 2.36 | 151.29 | NA | 0 | NA | 0.59 | 1.42
Results from laboratories 1.07 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 241b (Certified Value 3.31 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.14 | 3.1736 | 3.4118 | 3.2399 | 3.03 | 3.552 | 3.26 | 3.1373 | 3.08 | 3 | 3.28 | 3.64 | 3.27 | 3.3 | 3.36 | 3.433 | 3.28 | 3.274 | 0.808 | 2.9255 | 3.462 | 3.3 | 3.3568 | 3.29 | 2.69 | 3.33 | 3.1 |  |  | 1
| 1 | 2 | 3.17 | 3.2108 | 3.4149 | 3.3556 | 2.98 | 3.6003 | 3.31 | 3.2864 | 3.05 | 2.99 | 3.28 | 3.44 | 3.23 | 3.25 | 3.13 | 3.366 | 3.23 | 3.334 | 0.777 | 2.8664 | 3.443 | 3.18 | 3.3448 | 3.24 | 2.66 | 3.31 | 3.2 |  |  | #N/A
| 1 | 3 | 3.17 | 3.1643 | 3.3928 | 3.2186 | 2.9 | 3.6003 | 3.39 | 3.1995 | 3.04 | 3.01 | 3.41 | 3.42 | 3.19 | 3.3 | 3.21 | 3.401 | 3.26 | 3.245 | 0.98 | 2.8555 | 3.4545 | 3.36 | 3.4359 | 3.4 | 2.68 | 3.37 | 3.1 |  |  | 16
| 1 | 4 | 3.22 | 3.1595 | 3.3983 | 3.2489 | 2.99 | 3.4073 | 3.3 | 3.1322 | 3.07 | 2.96 | 3.42 | 3.52 | 3.19 | 3.26 | 3.4 | 3.344 | 3.3 | 3.285 | 0.87 | 2.8966 | 3.421 | 3.28 | 3.3218 | 3.35 | 2.63 | 3.34 | 3.2 |  |  | 3.3075
| 1 | 5 | 3.19 | 3.2025 | 3.4075 | 3.2401 | 2.92 | 3.5219 | 3.34 | 3.2448 | 3.03 | 2.98 | 3.24 | 3.6 | 3.57 | 3.29 | 3.34 | 3.335 | 3.36 | 3.223 | 0.4 | 2.852 | 3.443 | 3.35 | 3.4234 | 3.37 | 2.7 | 3.31 | 3.5 |  |  | 40
| 1 | 6 | 3.19 | 3.1797 | 3.3901 | 3.3105 | 3.19 | 3.534 | 3.41 | 3.2793 | 3.09 | 3 | 3.24 | 3.48 | 3.24 | 3.32 | 3.36 | 3.293 | 3.23 | 3.356 | 0.36 | 2.8669 | 3.4245 | 3.26 | 3.3207 | 3.4 | 2.76 | 3.35 | 3.5
| Mean |  | 3.18 | 3.1817 | 3.4026 | 3.2689 | 3.0017 | 3.536 | 3.335 | 3.2132 | 3.06 | 2.99 | 3.3117 | 3.5167 | 3.2817 | 3.2867 | 3.3 | 3.362 | 3.2767 | 3.2862 | 0.69917 | 2.8772 | 3.4413 | 3.2883 | 3.3673 | 3.3417 | 2.6867 | 3.335 | 3.2667
| Median |  | 3.18 | 3.1766 | 3.4029 | 3.2445 | 2.985 | 3.543 | 3.325 | 3.2222 | 3.06 | 2.995 | 3.28 | 3.5 | 3.235 | 3.295 | 3.35 | 3.355 | 3.27 | 3.2795 | 0.7925 | 2.8666 | 3.443 | 3.29 | 3.3508 | 3.36 | 2.685 | 3.335 | 3.2
| Std Dev. |  | 0.026833 | 0.020711 | 0.010283 | 0.052635 | 0.10381 | 0.071122 | 0.056833 | 0.06817 | 0.023664 | 0.017889 | 0.082077 | 0.088015 | 0.14456 | 0.026583 | 0.10564 | 0.049784 | 0.049261 | 0.050988 | 0.25709 | 0.028436 | 0.016142 | 0.065853 | 0.050414 | 0.064317 | 0.043665 | 0.023452 | 0.18619
| Rel.Std.Dev. |  | 0.8438% | 0.6509% | 0.3022% | 1.61% | 3.458% | 2.011% | 1.704% | 2.122% | 0.7733% | 0.5983% | 2.478% | 2.503% | 4.405% | 0.8088% | 3.201% | 1.481% | 1.503% | 1.552% | 36.77% | 0.9883% | 0.4691% | 2.003% | 1.497% | 1.925% | 1.625% | 0.7032% | 5.7%
| PDM3 |  | -3.854% | -3.801% | 2.875% | -1.165% | -9.246% | 6.909% | 0.8325% | -2.849% | -7.482% | -9.598% | 0.1271% | 6.325% | -0.78% | -0.6288% | -0.2257% | 1.649% | -0.9312% | -0.6439% | -78.86% | -13.01% | 4.047% | -0.5784% | 1.808% | 1.034% | -18.77% | 0.8325% | -1.233%
| Z-Score (Absolute) |  | 0.94 | 0.93 | 1.04 | 0.15 | 2.53 | 2.22 | 0.43 | 0.65 | 2.01 | 2.63 | 0.23 | 2.05 | 0.04 | 0 | 0.12 | 0.67 | 0.08 | 0 | 23 | 3.64 | 1.38 | 0.02 | 0.72 | 0.49 | 5.33 | 0.43 | 0.17
| Analytical results for Mn in OREAS 241b (Certified Value 0.133 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.13 | 0.1345 | 0.13515 | 0.13096 | 0.127 | 0.1487 | 0.128 | 0.1291 | 0.1227 | 0.1593 | 0.1255 | 0.1465 | 0.1355 | 0.1325 | 0.1334 | 0.1378 | 0.133 | 0.1337 | 0.069 | 0.12879 | 0.1335 | 0.127 | 0.1207 | 0.134 | 0.1322 | 0.1306 | 0.141 |  |  | 1
| 1 | 2 | 0.13 | 0.1358 | 0.14077 | 0.13698 | 0.124 | 0.1518 | 0.131 | 0.1361 | 0.1229 | 0.1599 | 0.126 | 0.1395 | 0.133 | 0.1315 | 0.127 | 0.135 | 0.1314 | 0.1363 | 0.07 | 0.12748 | 0.1325 | 0.126 | 0.12016 | 0.133 | 0.1331 | 0.1293 | 0.138 |  |  | 23
| 1 | 3 | 0.13 | 0.1341 | 0.1379 | 0.13022 | 0.123 | 0.151 | 0.134 | 0.1327 | 0.1231 | 0.1592 | 0.132 | 0.1385 | 0.1315 | 0.1325 | 0.1309 | 0.1362 | 0.1328 | 0.1327 | 0.078 | 0.12928 | 0.1335 | 0.1305 | 0.12297 | 0.139 | 0.1333 | 0.1319 | 0.136 |  |  | 16
| 1 | 4 | 0.14 | 0.1351 | 0.13855 | 0.13222 | 0.124 | 0.1425 | 0.13 | 0.1316 | 0.1227 | 0.1559 | 0.134 | 0.1415 | 0.1315 | 0.1305 | 0.1371 | 0.1367 | 0.1342 | 0.1346 | 0.07 | 0.12916 | 0.133 | 0.1275 | 0.11995 | 0.137 | 0.1319 | 0.1307 | 0.141 |  |  | 0.13298
| 1 | 5 | 0.13 | 0.1365 | 0.13588 | 0.13205 | 0.124 | 0.1479 | 0.132 | 0.1339 | 0.1228 | 0.1358 | 0.1265 | 0.145 | 0.147 | 0.131 | 0.1337 | 0.1339 | 0.1365 | 0.1319 | 0.069 | 0.12623 | 0.1325 | 0.131 | 0.12297 | 0.138 | 0.1366 | 0.1302 | 0.138 |  |  | 41
| 1 | 6 | 0.13 | 0.1338 | 0.13962 | 0.1345 | 0.135 | 0.1487 | 0.135 | 0.1354 | 0.1226 | 0.1348 | 0.1255 | 0.141 | 0.133 | 0.134 | 0.1353 | 0.1327 | 0.1319 | 0.1382 | 0.066 | 0.1275 | 0.1325 | 0.128 | 0.12027 | 0.14 | 0.1422 | 0.1311 | 0.136
| Mean |  | 0.13167 | 0.13497 | 0.13798 | 0.13282 | 0.12617 | 0.14843 | 0.13167 | 0.13313 | 0.1228 | 0.15082 | 0.12825 | 0.142 | 0.13525 | 0.132 | 0.1329 | 0.13538 | 0.1333 | 0.13457 | 0.070333 | 0.12807 | 0.13292 | 0.12833 | 0.12117 | 0.13683 | 0.13488 | 0.13063 | 0.13833
| Median |  | 0.13 | 0.1348 | 0.13822 | 0.13214 | 0.124 | 0.1487 | 0.1315 | 0.1333 | 0.12275 | 0.15755 | 0.12625 | 0.14125 | 0.133 | 0.132 | 0.13355 | 0.1356 | 0.1329 | 0.13415 | 0.0695 | 0.12815 | 0.13275 | 0.12775 | 0.12049 | 0.1375 | 0.1332 | 0.13065 | 0.138
| Std Dev. |  | 0.0040825 | 0.0010386 | 0.0021534 | 0.0025021 | 0.004535 | 0.0032727 | 0.002582 | 0.002582 | 0.00017889 | 0.012105 | 0.0037517 | 0.0031305 | 0.0059393 | 0.0012649 | 0.0035525 | 0.0018862 | 0.0018417 | 0.0023475 | 0.0040332 | 0.0012031 | 0.0004916 | 0.0019916 | 0.0014157 | 0.0027869 | 0.0039565 | 0.00087101 | 0.0022509
| Rel.Std.Dev. |  | 3.101% | 0.7695% | 1.561% | 1.884% | 3.594% | 2.205% | 1.961% | 1.939% | 0.1457% | 8.026% | 2.925% | 2.205% | 4.391% | 0.9583% | 2.673% | 1.393% | 1.382% | 1.744% | 5.734% | 0.9394% | 0.3699% | 1.552% | 1.168% | 2.037% | 2.933% | 0.6668% | 1.627%
| PDM3 |  | -0.9881% | 1.493% | 3.759% | -0.1209% | -5.124% | 11.62% | -0.9881% | 0.1148% | -7.656% | 13.41% | -3.557% | 6.782% | 1.707% | -0.7374% | -0.06062% | 1.807% | 0.2402% | 1.193% | -47.11% | -3.692% | -0.04808% | -3.495% | -8.88% | 2.897% | 1.431% | -1.765% | 4.025%
| Z-Score (Absolute) |  | 0.34 | 0.56 | 1.38 | 0.03 | 1.85 | 4.24 | 0.34 | 0.06 | 2.77 | 4.89 | 1.28 | 2.48 | 0.64 | 0.25 | 0 | 0.67 | 0.1 | 0.45 | 17.11 | 1.32 | 0 | 1.25 | 3.21 | 1.07 | 0.54 | 0.62 | 1.48
| Analytical results for Mo in OREAS 241b (Certified Value 2.19 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.51 | 2.3 | 2.36 | 2.2097 | 2.1449 | < 1 | 3.06 | 2.3 | 2 | 2 | 3.8 | 2.25 | 2.3 | 2.24 | 2.23 | < 2 | 2 | 1.9 | 2.1 | 2.4 | 3 | 2.2803 | 2.28 | 1.993 | 2.41 | 2.6 | 2.13 | 2.23 |  | 1
| 1 | 2 | 2.52 | 2.3 | 2 | 2.2452 | 2.171 | < 1 | 2.97 | 2.3 | 2.2 | 2 | 3.8 | 2.14 | 2.3 | 2.21 | 2.16 | < 2 | 2.1 | 2 | 2.15 | 4.1 | 2 | 2.2589 | 2.12 | 1.9725 | 2.36 | 2.83 | 2.11 | 2.19 |  | 24
| 1 | 3 | 2.52 | 2.3 | 3.62 | 2.2252 | 2.1205 | < 1 | 2.71 | 2.1 | 2.2 | 2 | 3.7 | 2.25 | 2.26 | 2.08 | 2.12 | < 2 | 2.1 | 1.9 | 2.1 | 2.68 | 3 | 2.2174 | 2.34 | 1.9663 | 2.29 | 2.57 | 2.15 | 2.11 |  | 16
| 1 | 4 | 2.52 | 2.2 | 3.07 | 2.2713 | 2.1667 | < 1 | 2.72 | 2.6 | 2.2 | 2 | 3.7 | 2.14 | 2.26 | 2.11 | 2.09 | 2 | 1.9 | 1.9 | 2.06 | 3.06 | 3 | 2.2406 | 2.19 | 2.0355 | 2.38 | 2.53 | 2.03 | 2.17 |  | 2.193
| 1 | 5 | 2.72 | 2.2 | 2.55 | 2.2479 | 2.091 | < 1 | 2.82 | 2.4 | 2.3 | 2 | 3 | 2.13 | 2.38 | 2.2 | 2.18 | 2 | 1.9 | 1.9 | 2.14 | 5.9 | 2 | 2.1973 | 2.32 | 1.9348 | 2.43 | 2.52 | 2.03 | 2.26 |  | 42
| 1 | 6 | 2.55 | 2.2 | 2.76 | 2.211 | 2.1557 | < 1 | 2.59 | 2.3 | 2.2 | 2 | 3.2 | 2.18 | 2.45 | 2.18 | 2.14 | < 2 | 2 | 1.9 | 2.14 | 5.74 | 3 | 2.3723 | 2.13 | 1.9333 | 2.39 | 2.84 | 2.04 | 2.33
| Mean |  | 2.5567 | 2.25 | 2.7267 | 2.2351 | 2.1416 |  | 2.8117 | 2.3333 | 2.1833 | 2 | 3.5333 | 2.1817 | 2.325 | 2.17 | 2.1533 | 2 | 2 | 1.9167 | 2.115 | 3.98 | 2.6667 | 2.2611 | 2.23 | 1.9726 | 2.3767 | 2.6483 | 2.0817 | 2.215
| Median |  | 2.52 | 2.25 | 2.655 | 2.2352 | 2.1503 |  | 2.77 | 2.3 | 2.2 | 2 | 3.7 | 2.16 | 2.3 | 2.19 | 2.15 | 2 | 2 | 1.9 | 2.12 | 3.58 | 3 | 2.2497 | 2.235 | 1.9694 | 2.385 | 2.585 | 2.075 | 2.21
| Std Dev. |  | 0.081158 | 0.054772 | 0.56758 | 0.024083 | 0.030667 |  | 0.17589 | 0.1633 | 0.098319 | 0 | 0.34448 | 0.055648 | 0.075299 | 0.061968 | 0.048854 | 0 | 0.089443 | 0.040825 | 0.034496 | 1.5383 | 0.5164 | 0.06186 | 0.096333 | 0.038433 | 0.048854 | 0.14743 | 0.054559 | 0.076354
| Rel.Std.Dev. |  | 3.174% | 2.434% | 20.82% | 1.078% | 1.432% |  | 6.256% | 6.999% | 4.503% | 0% | 9.749% | 2.551% | 3.239% | 2.856% | 2.269% | 0% | 4.472% | 2.13% | 1.631% | 38.65% | 19.36% | 2.736% | 4.32% | 1.948% | 2.056% | 5.567% | 2.621% | 3.447%
| PDM3 |  | 16.59% | 2.601% | 24.34% | 1.92% | -2.34% |  | 28.21% | 6.402% | -0.4386% | -8.799% | 61.12% | -0.5146% | 6.022% | -1.047% | -1.807% | -8.799% | -8.799% | -12.6% | -3.555% | 81.49% | 21.6% | 3.109% | 1.689% | -10.05% | 8.378% | 20.77% | -5.075% | 1.005%
| Z-Score (Absolute) |  | 1.92 | 0.12 | 2.91 | 0.03 | 0.52 | 10.14 | 3.41 | 0.61 | 0.27 | NA | 7.64 | 0.28 | 0.56 | 0.35 | 0.45 | NA | 1.35 | 1.84 | 0.67 | 10.26 | NA | 0.18 | 0 | 1.51 | 0.86 | 2.45 | 0.87 | 0.09
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 241b (Certified Value 1.93 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.91 | 1.8991 | 1.8287 | 1.7379 | 1.9318 | 1.86 | 2.144 | 1.84 | 1.7775 | 1.84 | 2.01 | 1.94 | 2.11 | 1.93 | 1.9 | 1.98 | 1.942 | 1.854 | 1.861 | 2.596 | 1.9802 | 2.01 | 1.93 | 2.416 | 1.93 | 2.27 | 2.06 | 1.97 |  | 1
| 1 | 2 | 1.93 | 1.9334 | 1.8502 | 1.807 | 1.9868 | 1.85 | 2.122 | 1.87 | 1.8654 | 1.85 | 1.98 | 1.93 | 2.01 | 1.91 | 1.86 | 1.86 | 1.918 | 1.866 | 1.897 | 3.157 | 1.9488 | 1.99 | 1.85 | 2.415 | 1.91 | 2.24 | 2.05 | 1.93 |  | #N/A
| 1 | 3 | 1.95 | 1.9023 | 1.7864 | 1.7757 | 1.9151 | 1.83 | 2.174 | 1.91 | 1.8327 | 1.84 | 1.97 | 2 | 1.99 | 1.89 | 1.9 | 1.92 | 1.928 | 1.883 | 1.844 | 2.946 | 1.9708 | 2.01 | 1.95 | 2.443 | 1.99 | 2.27 | 2.09 | 1.9 |  | 16
| 1 | 4 | 1.99 | 1.8881 | 2.0041 | 1.7806 | 1.9287 | 1.84 | 2.018 | 1.87 | 1.7849 | 1.84 | 1.94 | 2.01 | 2.05 | 1.89 | 1.87 | 2.01 | 1.869 | 1.88 | 1.866 | 3.462 | 1.9558 | 1.99 | 1.93 | 2.387 | 1.96 | 2.25 | 2.06 | 1.97 |  | 1.9265
| 1 | 5 | 1.97 | 1.9266 | 1.9354 | 1.8352 | 1.9327 | 1.79 | 2.077 | 1.89 | 1.8229 | 1.86 | 1.96 | 1.95 | 2.09 | 2.09 | 1.88 | 1.99 | 1.891 | 1.886 | 1.835 | 4.055 | 1.9278 | 1.98 | 1.97 | 2.459 | 1.97 | 2.29 | 2.05 | 1.89 |  | 43
| 1 | 6 | 1.94 | 1.8819 | 1.9413 | 1.8004 | 1.9637 | 1.94 | 2.107 | 1.92 | 1.833 | 1.86 | 1.96 | 1.94 | 2.02 | 1.91 | 1.9 | 2.03 | 1.864 | 1.879 | 1.911 | 4.115 | 1.9368 | 2 | 1.89 | 2.359 | 1.99 | 2.32 | 2.06 | 1.86
| Mean |  | 1.9483 | 1.9052 | 1.891 | 1.7895 | 1.9431 | 1.8517 | 2.107 | 1.8833 | 1.8194 | 1.8483 | 1.97 | 1.9617 | 2.045 | 1.9367 | 1.885 | 1.965 | 1.902 | 1.8747 | 1.869 | 3.3885 | 1.9534 | 1.9967 | 1.92 | 2.4132 | 1.9583 | 2.2733 | 2.0617 | 1.92
| Median |  | 1.945 | 1.9007 | 1.8928 | 1.7905 | 1.9323 | 1.845 | 2.1145 | 1.88 | 1.8278 | 1.845 | 1.965 | 1.945 | 2.035 | 1.91 | 1.89 | 1.985 | 1.9045 | 1.8795 | 1.8635 | 3.3095 | 1.9523 | 1.995 | 1.93 | 2.4155 | 1.965 | 2.27 | 2.06 | 1.915
| Std Dev. |  | 0.028577 | 0.02066 | 0.08223 | 0.033002 | 0.026692 | 0.049565 | 0.054597 | 0.029439 | 0.032973 | 0.0098319 | 0.023664 | 0.034303 | 0.047223 | 0.076594 | 0.017607 | 0.063482 | 0.032206 | 0.012226 | 0.029672 | 0.60908 | 0.019874 | 0.012111 | 0.043359 | 0.036389 | 0.032506 | 0.028752 | 0.01472 | 0.044721
| Rel.Std.Dev. |  | 1.467% | 1.084% | 4.348% | 1.844% | 1.374% | 2.677% | 2.591% | 1.563% | 1.812% | 0.5319% | 1.201% | 1.749% | 2.309% | 3.955% | 0.934% | 3.231% | 1.693% | 0.6522% | 1.588% | 17.97% | 1.017% | 0.6065% | 2.258% | 1.508% | 1.66% | 1.265% | 0.714% | 2.329%
| PDM3 |  | 1.135% | -1.102% | -1.841% | -7.112% | 0.865% | -3.883% | 9.371% | -2.239% | -5.558% | -4.056% | 2.26% | 1.827% | 6.153% | 0.5293% | -2.153% | 2% | -1.27% | -2.689% | -2.983% | 75.89% | 1.397% | 3.644% | -0.3359% | 25.26% | 1.654% | 18.01% | 7.018% | -0.3359%
| Z-Score (Absolute) |  | 0.1 | 0.42 | 0.59 | 1.82 | 0.04 | 1.07 | 2.02 | 0.68 | 1.46 | 1.11 | 0.36 | 0.26 | 1.27 | 0.04 | 0.66 | 0.3 | 0.46 | 0.79 | 0.86 | 17.53 | 0.16 | 0.69 | 0.24 | 5.73 | 0.22 | 4.03 | 1.47 | 0.24
| Analytical results for Nb in OREAS 241b (Certified Value 3.66 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.7 | 3.66 | 18.78 | 2.1798 | 3.8484 | 0.7 | 4.9 | 3.6 | 3.53 | < 5 | 3.7 | 3.8 | 3.9 | 3.6 | 3.4 | 24 | 3.3 | 3 | 3.4 | 5 | 3.2725 | 3.7 | 4.7765 | 4 | 3.6 | 3.7 | 3.6 |  |  | 1
| 1 | 2 | 3.6 | 3.8 | 19.06 | 2.2103 | 3.7984 | 0.8 | 4.8 | 3.6 | 3.6 | < 5 | 3.8 | 3.8 | 3.7 | 4 | 3.6 | 22 | 3.7 | 3 | 3.7 | 5 | 3.11 | 3.2 | 4.8086 | 3.8 | 3.7 | 3.6 | 3.4 |  |  | 7
| 1 | 3 | 3.7 | 3.89 | 20.86 | 2.1819 | 3.6771 | 0.4 | 4.4 | 3.6 | 3.61 | < 5 | 3.8 | 3.9 | 3.7 | 3.7 | 3.4 | 23 | 3.7 | 3.1 | 3.5 | 5 | 3.3407 | 3.8 | 4.8105 | 4 | 3.7 | 3.7 | 3.4 |  |  | 16
| 1 | 4 | 3.5 | 3.74 | 18.25 | 2.0605 | 3.6265 | 0.4 | 4.4 | 3.8 | 3.57 | < 5 | 3.8 | 3.9 | 4 | 3.6 | 3.4 | 24 | 3.4 | 3 | 3.6 | 5 | 3.2477 | 3.7 | 4.8218 | 3.9 | 3.6 | 3.4 | 3.5 |  |  | 3.6613
| 1 | 5 | 4.6 | 3.8 | 20.31 | 2.228 | 3.6725 | 0.8 | 4.4 | 3.7 | 3.64 | < 5 | 3.9 | 3.8 | 4.1 | 3.7 | 3.5 | 24 | 3.6 | 3.1 | 3.7 | 16 | 3.4096 | 3.8 | 4.6575 | 5.6 | 3.8 | 3.4 | 3.6 |  |  | 44
| 1 | 6 | 3.8 | 3.77 | 21.25 | 2.151 | 3.7741 | 0.6 | 4.1 | 3.7 | 3.71 | < 5 | 3.8 | 3.8 | 4 | 3.6 | 3.4 | 23 | 3.5 | 3 | 3.7 | 9 | 3.3678 | 3.5 | 4.7628 | 4.3 | 3.8 | 3.4 | 3.8
| Mean |  | 3.8167 | 3.7767 | 19.7517 | 2.1686 | 3.7328 | 0.61667 | 4.5 | 3.6667 | 3.61 |  | 3.8 | 3.8333 | 3.9 | 3.7 | 3.45 | 23.3333 | 3.5333 | 3.0333 | 3.6 | 7.5 | 3.2914 | 3.6167 | 4.773 | 4.2667 | 3.7 | 3.5333 | 3.55
| Median |  | 3.7 | 3.785 | 19.685 | 2.1809 | 3.7256 | 0.65 | 4.4 | 3.65 | 3.605 |  | 3.8 | 3.8 | 3.95 | 3.65 | 3.4 | 23.5 | 3.55 | 3 | 3.65 | 5 | 3.3066 | 3.7 | 4.7926 | 4 | 3.7 | 3.5 | 3.55
| Std Dev. |  | 0.39707 | 0.07607 | 1.2217 | 0.059283 | 0.086468 | 0.18348 | 0.29665 | 0.08165 | 0.061644 |  | 0.063246 | 0.05164 | 0.16733 | 0.15492 | 0.083666 | 0.8165 | 0.1633 | 0.05164 | 0.12649 | 4.4609 | 0.10716 | 0.23166 | 0.060847 | 0.67429 | 0.089443 | 0.15055 | 0.15166
| Rel.Std.Dev. |  | 10.4% | 2.014% | 6.185% | 2.734% | 2.316% | 29.75% | 6.592% | 2.227% | 1.708% |  | 1.664% | 1.347% | 4.291% | 4.187% | 2.425% | 3.499% | 4.622% | 1.702% | 3.514% | 59.48% | 3.256% | 6.405% | 1.275% | 15.8% | 2.417% | 4.261% | 4.272%
| PDM3 |  | 4.245% | 3.152% | 439.5% | -40.77% | 1.955% | -83.16% | 22.91% | 0.1479% | -1.4% |  | 3.79% | 4.7% | 6.521% | 1.058% | -5.77% | 537.3% | -3.494% | -17.15% | -1.673% | 104.8% | -10.1% | -1.218% | 30.36% | 16.54% | 1.058% | -3.494% | -3.039%
| Z-Score (Absolute) |  | 0.6 | 0.42 | 72.23 | 6.81 | 0.22 | 13.79 | 3.67 | 0.07 | 0.33 | 5.32 | 0.52 | 0.67 | 0.97 | 0.07 | 1.05 | 88.33 | 0.67 | 2.92 | 0.37 | NA | 1.76 | 0.3 | 4.9 | 2.62 | 0.07 | 0.67 | 0.6
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Nd in OREAS 241b (Certified Value 9.2 ppm) | [30 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 9.13 | 5.25 | 8.3598 | 8.8 | 11.3 | 8.9 | 10 | 9.1271 | [21 empty] | 1
| 1 | 2 | 9.15 | 5.37 | 8.395 | 8.3 | 10.9 | 8.7 | 10 | 9.4629 | [21 empty] | 26
| 1 | 3 | 9.14 | 5.31 | 8.3746 | 8.4 | 9.8 | 8.7 | 10 | 9.4852 | [21 empty] | 16
| 1 | 4 | 9.38 | 5.03 | 8.3147 | 8.7 | 10 | 8.7 | 10 | 9.3851 | [21 empty] | 9.2007
| 1 | 5 | 9.15 | 5.18 | 8.3998 | 8.2 | 9.8 | 8.7 | 9.5 | 9.3067 | [21 empty] | 45
| 1 | 6 | 9.03 | 5.09 | 8.3717 | 9 | 9.4 | 8.8 | 10.5 | 9.3682
| Mean |  | 9.1633 | 5.205 | 8.3693 | 8.5667 | 10.2 | 8.75 | 10 | 9.3559
| Median |  | 9.145 | 5.215 | 8.3731 | 8.55 | 9.9 | 8.7 | 10 | 9.3766
| Std Dev. |  | 0.11553 | 0.13019 | 0.030645 | 0.31411 | 0.73485 | 0.083666 | 0.31623 | 0.12956
| Rel.Std.Dev. |  | 1.261% | 2.501% | 0.3662% | 3.667% | 7.204% | 0.9562% | 3.162% | 1.385%
| PDM3 |  | -0.4065% | -43.43% | -9.037% | -6.891% | 10.86% | -4.899% | 8.687% | 1.686%
| Z-Score (Absolute) |  | 0.28 | 5.13 | 0.8 | 0.53 | 1.7 | 0.28 | 1.43 | 0.55
| Analytical results for Ni in OREAS 241b (Certified Value 77 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
1
| 1 | 1 | 76.7 | 78 | 75.48 | 76.57 | 82.2234 | 79 | 110 | 76 | 76 | 72 | 77.6 | 82.8 | 85.7 | 78.1 | 78.8 | 71 | 78 | 75.9 | 76 | 63.58 | 76.1112 | 80 | 81.0731 | 77.8 | 60.248 | 70 | 80.5 | 76.8 | 81.8 | 1
| 1 | 2 | 77.2 | 78.4 | 76.68 | 78.47 | 85.1963 | 78 | 108 | 77 | 79 | 73 | 77 | 83.4 | 80.9 | 78.8 | 80 | 68 | 76 | 74 | 78 | 63.08 | 75.9531 | 75 | 78.5657 | 71.3 | 59.855 | 70.8 | 80.7 | 78.8 | 80 | 27
| 1 | 3 | 76.9 | 77.9 | 74.84 | 78.06 | 81.7799 | 78 | 99.1 | 78 | 78.6 | 71 | 76.9 | 84.8 | 78.4 | 77.1 | 79.1 | 69 | 77 | 76.1 | 75 | 63.57 | 77.5341 | 75 | 79.8334 | 81 | 59.6893 | 72.7 | 82.4 | 77.5 | 78.7 | 16
| 1 | 4 | 79.1 | 78.1 | 74.79 | 77.65 | 82.7537 | 78 | 98.4 | 77 | 75.8 | 73 | 75.7 | 83.4 | 84.7 | 76.4 | 77.3 | 72 | 78 | 76.6 | 76 | 61.78 | 76.6134 | 80 | 79.6645 | 78.3 | 62.6028 | 69.8 | 80.8 | 75.2 | 80.7 | 77.4656
| 1 | 5 | 76.7 | 79.3 | 74.68 | 78.39 | 82.9764 | 79 | 97.2 | 78 | 78.4 | 75 | 74.8 | 82.9 | 88.4 | 81.5 | 79 | 71 | 76 | 77.4 | 74 | 63.66 | 76.1949 | 80 | 78.735 | 80.7 | 58.1427 | 71.8 | 84.6 | 78.2 | 79.2 | 46
| 1 | 6 | 76.7 | 78 | 76.52 | 78.54 | 84.5632 | 78 | 90.3 | 79 | 78.7 | 72 | 75.8 | 83.4 | 85.3 | 77.5 | 79.8 | 72 | 76 | 76.1 | 79 | 63.82 | 76.0275 | 75 | 78.889 | 72 | 59.4251 | 73.1 | 85.6 | 76.3 | 77.9
| Mean |  | 77.2167 | 78.2833 | 75.4983 | 77.9467 | 83.2488 | 78.3333 | 100.5 | 77.5 | 77.75 | 72.6667 | 76.3 | 83.45 | 83.9 | 78.2333 | 79 | 70.5 | 76.8333 | 76.0167 | 76.3333 | 63.2483 | 76.4057 | 77.5 | 79.4601 | 76.85 | 59.9938 | 71.3667 | 82.4333 | 77.1333 | 79.7167
| Median |  | 76.8 | 78.05 | 75.16 | 78.225 | 82.865 | 78 | 98.75 | 77.5 | 78.5 | 72.5 | 76.35 | 83.4 | 85 | 77.8 | 79.05 | 71 | 76.5 | 76.1 | 76 | 63.575 | 76.1531 | 77.5 | 79.2768 | 78.05 | 59.7721 | 71.3 | 81.6 | 77.15 | 79.6
| Std Dev. |  | 0.94322 | 0.52694 | 0.89948 | 0.7508 | 1.3453 | 0.5164 | 7.3185 | 1.0488 | 1.4474 | 1.3663 | 1.0392 | 0.71484 | 3.615 | 1.7996 | 0.95708 | 1.6432 | 0.98319 | 1.1268 | 1.8619 | 0.7608 | 0.59925 | 2.7386 | 0.94179 | 4.2279 | 1.4652 | 1.3866 | 2.197 | 1.311 | 1.4134
| Rel.Std.Dev. |  | 1.222% | 0.6731% | 1.191% | 0.9632% | 1.616% | 0.6592% | 7.282% | 1.353% | 1.862% | 1.88% | 1.362% | 0.8566% | 4.309% | 2.3% | 1.211% | 2.331% | 1.28% | 1.482% | 2.439% | 1.203% | 0.7843% | 3.534% | 1.185% | 5.501% | 2.442% | 1.943% | 2.665% | 1.7% | 1.773%
| PDM3 |  | -0.3213% | 1.056% | -2.54% | 0.621% | 7.466% | 1.12% | 29.74% | 0.04441% | 0.3671% | -6.195% | -1.505% | 7.725% | 8.306% | 0.9911% | 1.981% | -8.992% | -0.8162% | -1.87% | -1.462% | -18.35% | -1.368% | 0.04441% | 2.575% | -0.7947% | -22.55% | -7.873% | 6.413% | -0.4289% | 2.906%
| Z-Score (Absolute) |  | 0.13 | 0.36 | 0.91 | 0.2 | 2.61 | 0.38 | 10.46 | 0 | 0.11 | 2.2 | 0.55 | 2.7 | 2.91 | 0.33 | 0.68 | 3.18 | 0.3 | 0.67 | 0.53 | 6.48 | 0.5 | 0 | 0.89 | 0.3 | 7.96 | 2.79 | 2.24 | 0.17 | 1.01
| Analytical results for P in OREAS 241b (Certified Value 0.044 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.0433 | 0.0435 | 0.042146 | 0.044183 | 0.039 | 0.0537 | 0.046 | 0.0394 | 0.0429 | 0.046 | 0.044 | 0.048 | 0.045 | 0.043 | 0.0429 | 0.046 | 0.042 | 0.044 | 0.044 | 0.045 | 0.043704 | 0.044 | 0.0447 | 0.0439 | 0.0444 |  |  |  |  | 1
| 1 | 2 | 0.0433 | 0.0466 | 0.042335 | 0.04556 | 0.041 | 0.0528 | 0.046 | 0.0425 | 0.0429 | 0.047 | 0.043 | 0.046 | 0.045 | 0.043 | 0.0408 | 0.046 | 0.042 | 0.046 | 0.044 | 0.043 | 0.043734 | 0.0433 | 0.0437 | 0.0435 | 0.0438 |  |  |  |  | #N/A
| 1 | 3 | 0.044 | 0.0458 | 0.042731 | 0.043922 | 0.039 | 0.0532 | 0.047 | 0.0434 | 0.0428 | 0.046 | 0.045 | 0.045 | 0.044 | 0.044 | 0.0416 | 0.046 | 0.044 | 0.044 | 0.0435 | 0.047 | 0.044549 | 0.0436 | 0.044 | 0.0443 | 0.0428 |  |  |  |  | 16
| 1 | 4 | 0.0438 | 0.0441 | 0.041997 | 0.044083 | 0.038 | 0.0502 | 0.046 | 0.0427 | 0.0434 | 0.045 | 0.046 | 0.047 | 0.044 | 0.044 | 0.0445 | 0.044 | 0.042 | 0.045 | 0.0435 | 0.045 | 0.043852 | 0.0429 | 0.0435 | 0.0439 | 0.0437 |  |  |  |  | 0.044165
| 1 | 5 | 0.0438 | 0.0439 | 0.042566 | 0.044032 | 0.041 | 0.0528 | 0.047 | 0.0438 | 0.0432 | 0.045 | 0.043 | 0.048 | 0.049 | 0.044 | 0.044 | 0.045 | 0.043 | 0.044 | 0.0435 | 0.046 | 0.044582 | 0.0437 | 0.0446 | 0.0436 | 0.0435 |  |  |  |  | 47
| 1 | 6 | 0.0439 | 0.0441 | 0.042021 | 0.045107 | 0.038 | 0.0541 | 0.047 | 0.0488 | 0.0431 | 0.046 | 0.044 | 0.046 | 0.044 | 0.044 | 0.0444 | 0.044 | 0.042 | 0.047 | 0.0425 | 0.043 | 0.043594 | 0.0449 | 0.0454 | 0.0439 | 0.0432
| Mean |  | 0.043683 | 0.044667 | 0.042299 | 0.044481 | 0.039333 | 0.0528 | 0.0465 | 0.043433 | 0.04305 | 0.045833 | 0.044167 | 0.046667 | 0.045167 | 0.043667 | 0.043033 | 0.045167 | 0.0425 | 0.045 | 0.0435 | 0.044833 | 0.044002 | 0.043733 | 0.044317 | 0.04385 | 0.043567
| Median |  | 0.0438 | 0.0441 | 0.042241 | 0.044133 | 0.039 | 0.053 | 0.0465 | 0.04305 | 0.043 | 0.046 | 0.044 | 0.0465 | 0.0445 | 0.044 | 0.04345 | 0.0455 | 0.042 | 0.0445 | 0.0435 | 0.045 | 0.043793 | 0.04365 | 0.0443 | 0.0439 | 0.0436
| Std Dev. |  | 0.00030605 | 0.001234 | 0.00030037 | 0.00068099 | 0.0013663 | 0.0013726 | 0.00054772 | 0.0030533 | 0.00022583 | 0.00075277 | 0.001169 | 0.0012111 | 0.0019408 | 0.0005164 | 0.0015501 | 0.00098319 | 0.00083666 | 0.0012649 | 0.00054772 | 0.0016021 | 0.000444 | 0.00068313 | 0.00071391 | 0.00028107 | 0.0005465
| Rel.Std.Dev. |  | 0.7006% | 2.763% | 0.7101% | 1.531% | 3.474% | 2.6% | 1.178% | 7.03% | 0.5246% | 1.642% | 2.647% | 2.595% | 4.297% | 1.183% | 3.602% | 2.177% | 1.969% | 2.811% | 1.259% | 3.573% | 1.009% | 1.562% | 1.611% | 0.641% | 1.254%
| PDM3 |  | -1.091% | 1.135% | -4.225% | 0.7156% | -10.94% | 19.55% | 5.286% | -1.657% | -2.525% | 3.777% | 0.003236% | 5.664% | 2.267% | -1.129% | -2.563% | 2.267% | -3.77% | 1.89% | -1.506% | 1.513% | -0.3688% | -0.9779% | 0.3429% | -0.7138% | -1.355%
| Z-Score (Absolute) |  | 0.26 | 0.54 | 1.38 | 0.39 | 3.79 | 7.14 | 2.03 | 0.46 | 0.77 | 1.49 | 0.13 | 2.16 | 0.94 | 0.27 | 0.79 | 0.94 | 1.22 | 0.81 | 0.41 | 0.67 | 0 | 0.22 | 0.26 | 0.12 | 0.35
| Analytical results for Pb in OREAS 241b (Certified Value 37.3 ppm) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 1
2
| 1 | 1 | 35.9 | 38.5 | 30.15 | 35.793 | 38.187 | 31 | 52.5 | 40 | 36.2 | 40 | 36.5 | 37.5 | 38.9 | 36.6 | 35.8 | 30 | 32 | 33.9 | 40.1 | 33.66 | 36.1524 | 44 | 37.4502 | 37.2 | 152.5676 | 38 | 37.6 | 37.3 | 39.9 | 1
| 1 | 2 | 36.2 | 38.9 | 32.02 | 35.271 | 39.1708 | 32 | 49.2 | 40 | 37.6 | 39 | 35 | 37.4 | 37.5 | 37 | 37 | 29 | 31 | 33.4 | 39.4 | 38.59 | 36.266 | 42 | 37.343 | 37.9 | 143.1211 | 37.5 | 37.1 | 37.8 | 38.7 | 28
| 1 | 3 | 36.3 | 38.7 | 30.72 | 35.98 | 39.052 | 32 | 45.7 | 39 | 37 | 40 | 35.6 | 36.6 | 36.8 | 36.7 | 37 | 30 | 34 | 33.5 | 38.9 | 37.01 | 37.4778 | 44 | 38.3246 | 37.5 | 139.465 | 37.8 | 37.2 | 37.6 | 37.5 | 16
| 1 | 4 | 36.7 | 38.1 | 29.92 | 35.7366 | 38.4396 | 32 | 45.1 | 37 | 36.9 | 39 | 35.9 | 36.7 | 38.5 | 35.6 | 35.8 | 30 | 36 | 34.1 | 39.9 | 33.42 | 36.9138 | 44 | 37.6675 | 36.6 | 140.8842 | 37.6 | 37.4 | 36.9 | 38.8 | 37.3356
| 1 | 5 | 36.6 | 38.6 | 32.08 | 35.7075 | 37.9625 | 32 | 44.5 | 38 | 37.4 | 40 | 35.5 | 37.4 | 39.9 | 37.7 | 35.6 | 31 | 39 | 35.4 | 38.9 | 34.69 | 35.7106 | 40 | 38.7387 | 38 | 141.3747 | 37.5 | 38.8 | 38.1 | 40.3 | 48
| 1 | 6 | 36.5 | 37.6 | 32.23 | 35.5814 | 39.267 | 31 | 43 | 39 | 37.5 | 39 | 36.2 | 37.2 | 39.7 | 37.3 | 36 | 31 | 37 | 34.2 | 39.4 | 48.06 | 36.4156 | 43 | 38.0964 | 41.7 | 149.2509 | 38.1 | 40.6 | 37.3 | 41.5
| Mean |  | 36.3667 | 38.4 | 31.1867 | 35.6782 | 38.6798 | 31.6667 | 46.6667 | 38.8333 | 37.1 | 39.5 | 35.7833 | 37.1333 | 38.55 | 36.8167 | 36.2 | 30.1667 | 34.8333 | 34.0833 | 39.4333 | 37.5717 | 36.4894 | 42.8333 | 37.9367 | 38.15 | 144.4439 | 37.75 | 38.1167 | 37.5 | 39.45
| Median |  | 36.4 | 38.55 | 31.37 | 35.7221 | 38.7458 | 32 | 45.4 | 39 | 37.2 | 39.5 | 35.75 | 37.3 | 38.7 | 36.85 | 35.9 | 30 | 35 | 34 | 39.4 | 35.85 | 36.3408 | 43.5 | 37.8819 | 37.7 | 142.2479 | 37.7 | 37.5 | 37.45 | 39.35
| Std Dev. |  | 0.29439 | 0.47329 | 1.0467 | 0.23821 | 0.55487 | 0.5164 | 3.5206 | 1.169 | 0.52154 | 0.54772 | 0.53448 | 0.38816 | 1.2194 | 0.71949 | 0.63246 | 0.75277 | 3.0605 | 0.71949 | 0.49666 | 5.5172 | 0.62169 | 1.6021 | 0.54419 | 1.8119 | 5.2485 | 0.25884 | 1.363 | 0.42426 | 1.4082
| Rel.Std.Dev. |  | 0.8095% | 1.233% | 3.356% | 0.6677% | 1.435% | 1.631% | 7.544% | 3.01% | 1.406% | 1.387% | 1.494% | 1.045% | 3.163% | 1.954% | 1.747% | 2.495% | 8.786% | 2.111% | 1.259% | 14.68% | 1.704% | 3.74% | 1.434% | 4.749% | 3.634% | 0.6857% | 3.576% | 1.131% | 3.57%
| PDM3 |  | -2.595% | 2.851% | -16.47% | -4.439% | 3.6% | -15.18% | 24.99% | 4.012% | -0.631% | 5.797% | -4.158% | -0.5417% | 3.253% | -1.39% | -3.042% | -19.2% | -6.702% | -8.711% | 5.619% | 0.6324% | -2.266% | 14.73% | 1.61% | 2.181% | 286.9% | 1.11% | 2.092% | 0.4404% | 5.663%
| Z-Score (Absolute) |  | 0.64 | 0.44 | 3.41 | 1.01 | 0.59 | 3.16 | 4.86 | 0.67 | 0.25 | 1.03 | 0.96 | 0.23 | 0.52 | 0.4 | 0.73 | 3.96 | 1.46 | 1.86 | 0.99 | 0 | 0.58 | 2.81 | 0.2 | 0.31 | 57.12 | 0.1 | 0.29 | 0.04 | 1
| Analytical results for Pr in OREAS 241b (Certified Value 1.99 ppm) | [30 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
2
| 1 | 1 | 2.04 | 1.72 | 1.7321 | 1.8 | 2.6 | 2.01 | 2.1 | 2.0115 | [21 empty] | 1
| 1 | 2 | 2 | 1.79 | 1.8359 | 1.8 | 2.6 | 1.98 | 2.2 | 2.0195 | [21 empty] | 29
| 1 | 3 | 2.03 | 1.77 | 1.7543 | 1.8 | 2.3 | 1.93 | 2.1 | 2.0527 | [21 empty] | 16
| 1 | 4 | 1.97 | 1.74 | 1.7949 | 1.9 | 2.3 | 2.02 | 2.2 | 2.0502 | [21 empty] | 1.9902
| 1 | 5 | 2.02 | 1.79 | 1.8295 | 1.8 | 2.3 | 1.95 | 2.1 | 2.0163 | [21 empty] | 49
| 1 | 6 | 2.02 | 1.76 | 1.7799 | 1.9 | 2.2 | 1.98 | 2.1 | 2.0335
| Mean |  | 2.0133 | 1.7617 | 1.7878 | 1.8333 | 2.3833 | 1.9783 | 2.1333 | 2.0306
| Median |  | 2.02 | 1.765 | 1.7874 | 1.8 | 2.3 | 1.98 | 2.1 | 2.0265
| Std Dev. |  | 0.025033 | 0.027869 | 0.040984 | 0.05164 | 0.17224 | 0.034303 | 0.05164 | 0.017747
| Rel.Std.Dev. |  | 1.243% | 1.582% | 2.292% | 2.817% | 7.227% | 1.734% | 2.421% | 0.874%
| PDM3 |  | 1.161% | -11.48% | -10.17% | -7.883% | 19.75% | -0.5972% | 7.191% | 2.031%
| Z-Score (Absolute) |  | 0.08 | 1.05 | 0.94 | 0.73 | 1.74 | 0.08 | 0.62 | 0.16
| Analytical results for Rb in OREAS 241b (Certified Value 18.1 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 18.6 | 17.54 | 16.5945 | 17.532 | 17.6 | 21.9 | 18.2 | 17.39 | 16.8 | 19.2 | 20 | 17.4 | 17.7 | 17.3 | 17.5 | 18.2 | 20 | 19.665 | 18 | 17.6952 | 19.8 | 17.8 | 18.2 | 18.5 | [5 empty] | 1
| 1 | 2 | 18.6 | 17.76 | 16.7194 | 17.6408 | 17 | 21.4 | 17.7 | 17.83 | 17.2 | 18.6 | 18.6 | 17.4 | 18 | 21.1 | 17.1 | 18.2 | 20.5 | 19.1172 | 16.1 | 17.6604 | 19.5 | 17.9 | 18.5 | 17.9 | [5 empty] | 30
| 1 | 3 | 18.4 | 17.48 | 16.5453 | 17.0543 | 17.5 | 19.5 | 18.7 | 18.01 | 17.6 | 19 | 18.7 | 17 | 17.8 | 19.4 | 17.4 | 17.8 | 20.5 | 19.0013 | 18.7 | 17.71 | 19.7 | 17.8 | 18 | 17.8 | [5 empty] | 16
| 1 | 4 | 17.9 | 17.44 | 16.5916 | 16.9691 | 17.5 | 19.6 | 17.7 | 17.56 | 17.9 | 18.6 | 19.9 | 16.8 | 17.2 | 17.9 | 17.8 | 17.9 | 21 | 18.256 | 17.9 | 17.6456 | 19.6 | 18.1 | 17.7 | 17.9 | [5 empty] | 18.0996
| 1 | 5 | 20.1 | 17.54 | 16.8216 | 17.1041 | 17.5 | 19.4 | 18.7 | 17.86 | 17.6 | 18.6 | 20.3 | 17.4 | 17.6 | 19.5 | 17.8 | 17.7 | 19.5 | 18.3551 | 19.2 | 17.2274 | 20.2 | 18.5 | 18.6 | 18.7 | [5 empty] | 50
| 1 | 6 | 19 | 17.15 | 16.5697 | 17.5963 | 18.6 | 18.2 | 18.6 | 17.86 | 18.1 | 18.6 | 20.2 | 17.2 | 16.9 | 20 | 17.3 | 18.5 | 20.5 | 18.3601 | 16.5 | 17.6267 | 19.8 | 19.3 | 17.7 | 19.7
| Mean |  | 18.7667 | 17.485 | 16.6404 | 17.3161 | 17.6167 | 20 | 18.2667 | 17.7517 | 17.5333 | 18.7667 | 19.6167 | 17.2 | 17.5333 | 19.2 | 17.4833 | 18.05 | 20.3333 | 18.7924 | 17.7333 | 17.5942 | 19.7667 | 18.2333 | 18.1167 | 18.4167
| Median |  | 18.6 | 17.51 | 16.5931 | 17.3181 | 17.5 | 19.55 | 18.4 | 17.845 | 17.6 | 18.6 | 19.95 | 17.3 | 17.65 | 19.45 | 17.45 | 18.05 | 20.5 | 18.6807 | 17.95 | 17.653 | 19.75 | 18 | 18.1 | 18.2
| Std Dev. |  | 0.74476 | 0.19796 | 0.10732 | 0.30477 | 0.52694 | 1.3842 | 0.4761 | 0.22973 | 0.47188 | 0.26583 | 0.76267 | 0.25298 | 0.40825 | 1.3914 | 0.27869 | 0.30166 | 0.5164 | 0.56151 | 1.2144 | 0.18234 | 0.24221 | 0.58538 | 0.38687 | 0.72778
| Rel.Std.Dev. |  | 3.969% | 1.132% | 0.6449% | 1.76% | 2.991% | 6.921% | 2.606% | 1.294% | 2.691% | 1.417% | 3.888% | 1.471% | 2.328% | 7.247% | 1.594% | 1.671% | 2.54% | 2.988% | 6.848% | 1.036% | 1.225% | 3.21% | 2.135% | 3.952%
| PDM3 |  | 3.685% | -3.396% | -8.062% | -4.329% | -2.668% | 10.5% | 0.9229% | -1.922% | -3.129% | 3.685% | 8.382% | -4.97% | -3.129% | 6.08% | -3.405% | -0.2742% | 12.34% | 3.828% | -2.024% | -2.792% | 9.21% | 0.7387% | 0.09414% | 1.752%
| Z-Score (Absolute) |  | 0.77 | 0.67 | 1.62 | 0.86 | 0.53 | 2.16 | 0.21 | 0.37 | 0.62 | 0.77 | 1.73 | 0.99 | 0.62 | 1.26 | 0.68 | 0.04 | 2.53 | 0.8 | 0.39 | 0.55 | 1.89 | 0.17 | 0.04 | 0.38
| Analytical results for Re in OREAS 241b (Certified Value 0.003 ppm) | [30 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.29 | 1.30 | 1.31 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | [10 empty] | 3
3
| 1 | 1 | < 0.002 | < 0.01 | 0.003 | < 0.05 | 0.004 | < 0.005 | 0.003 | 0.003 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0156 | 0.006 | 0.003 | 0.005 | [10 empty] | 1
| 1 | 2 | 0.002 | < 0.01 | 0.003 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.003 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0162 | 0.007 | 0.002 | 0.005 | [10 empty] | 31
| 1 | 3 | < 0.002 | < 0.01 | 0.043 | < 0.05 | < 0.002 | < 0.005 | 0.003 | 0.003 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0166 | 0.007 | 0.003 | 0.006 | [10 empty] | 16
| 1 | 4 | 0.003 | < 0.01 | 0.003 | < 0.05 | < 0.002 | < 0.005 | 0.002 | 0.003 | 0.002 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.003 | 0.0151 | 0.008 | 0.002 | 0.005 | [10 empty] | 0.0025963
| 1 | 5 | 0.003 | < 0.01 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.003 | 0.003 | 0.002 | < 0.002 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0163 | 0.004 | 0.003 | 0.004 | [10 empty] | 51
| 1 | 6 | < 0.002 | < 0.01 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.003 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.005 | < 0.05 | 0.002 | 0.0153 | 0.007 | 0.002 | 0.005
| Mean |  | 0.0026667 |  | 0.0093333 |  | 0.003 |  | 0.0026667 | 0.003 | 0.0021667 | 0.0026 |  |  |  |  | 0.0021667 | 0.01585 | 0.0065 | 0.0025 | 0.005
| Median |  | 0.003 |  | 0.003 |  | 0.003 |  | 0.003 | 0.003 | 0.002 | 0.003 |  |  |  |  | 0.002 | 0.0159 | 0.007 | 0.0025 | 0.005
| Std Dev. |  | 0.00057735 |  | 0.016501 |  | 0.001 |  | 0.0005164 | 4.7507e-19 | 0.00040825 | 0.00054772 |  |  |  |  | 0.00040825 | 0.00060249 | 0.0013784 | 0.00054772 | 0.00063246
| Rel.Std.Dev. |  | 21.65% |  | 176.8% |  | 33.33% |  | 19.36% | 1.584e-14% | 18.84% | 21.07% |  |  |  |  | 18.84% | 3.801% | 21.21% | 21.91% | 12.65%
| PDM3 |  | 2.71% |  | 259.5% |  | 15.55% |  | 2.71% | 15.55% | -16.55% | 0.1427% |  |  |  |  | -16.55% | 510.5% | 150.4% | -3.709% | 92.58%
| Z-Score (Absolute) |  | 0.67 | 1.16 | 3.66 | 12.72 | 0.58 | 0.29 | 0.19 | 0 | 0.48 | 0.39 | 1443.21 | 12.72 | 0.29 | 12.72 | 0.48 | 7.43 | 2.02 | 0.29 | 1.16
| Analytical results for S in OREAS 241b (Certified Value 0.496 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.49 | 0.5 | 0.51376 | 0.51082 | 0.47 | 0.555 | 0.481 | 0.47 | 0.49 | 0.48 | 0.49 | 0.55 | 0.49 | 0.49 | 0.48 | 0.509 | 0.4 | 0.51 | 0.48987 | 0.505 | 0.48 | 0.43855 | 0.51 | 0.49 | 0.5 | 0.51 |  |  |  | 1
| 1 | 2 | 0.49 | 0.5 | 0.52009 | 0.52756 | 0.48 | 0.562 | 0.487 | 0.49 | 0.49 | 0.48 | 0.48 | 0.52 | 0.49 | 0.49 | 0.45 | 0.501 | 0.4 | 0.535 | 0.47939 | 0.5 | 0.45 | 0.44587 | 0.5 | 0.5 | 0.5 | 0.51 |  |  |  | 16
| 1 | 3 | 0.49 | 0.49 | 0.51986 | 0.49855 | 0.47 | 0.554 | 0.495 | 0.48 | 0.49 | 0.49 | 0.5 | 0.51 | 0.48 | 0.5 | 0.46 | 0.508 | 0.4 | 0.517 | 0.4837 | 0.5 | 0.48 | 0.47306 | 0.52 | 0.5 | 0.51 | 0.49 |  |  |  | 16
| 1 | 4 | 0.52 | 0.51 | 0.51971 | 0.50032 | 0.48 | 0.541 | 0.482 | 0.48 | 0.5 | 0.5 | 0.51 | 0.53 | 0.48 | 0.49 | 0.49 | 0.497 | 0.4 | 0.52 | 0.47931 | 0.5 | 0.48 | 0.45942 | 0.52 | 0.49 | 0.5 | 0.51 |  |  |  | 0.49638
| 1 | 5 | 0.49 | 0.51 | 0.52257 | 0.50638 | 0.49 | 0.551 | 0.493 | 0.47 | 0.49 | 0.5 | 0.49 | 0.54 | 0.53 | 0.49 | 0.48 | 0.5 | 0.5 | 0.507 | 0.47926 | 0.505 | 0.49 | 0.43509 | 0.5 | 0.5 | 0.5 | 0.47 |  |  |  | 52
| 1 | 6 | 0.49 | 0.48 | 0.51425 | 0.5108 | 0.48 | 0.554 | 0.499 | 0.51 | 0.49 | 0.49 | 0.49 | 0.52 | 0.48 | 0.5 | 0.49 | 0.492 | 0.4 | 0.542 | 0.48128 | 0.5 | 0.47 | 0.45839 | 0.51 | 0.52 | 0.51 | 0.5
| Mean |  | 0.495 | 0.49833 | 0.51837 | 0.50907 | 0.47833 | 0.55283 | 0.4895 | 0.48333 | 0.49167 | 0.49 | 0.49333 | 0.52833 | 0.49167 | 0.49333 | 0.475 | 0.50117 | 0.41667 | 0.52183 | 0.48213 | 0.50167 | 0.475 | 0.45173 | 0.51 | 0.5 | 0.50333 | 0.49833
| Median |  | 0.49 | 0.5 | 0.51979 | 0.50859 | 0.48 | 0.554 | 0.49 | 0.48 | 0.49 | 0.49 | 0.49 | 0.525 | 0.485 | 0.49 | 0.48 | 0.5005 | 0.4 | 0.5185 | 0.48033 | 0.5 | 0.48 | 0.45213 | 0.51 | 0.5 | 0.5 | 0.505
| Std Dev. |  | 0.012247 | 0.01169 | 0.0035457 | 0.010413 | 0.0075277 | 0.0068532 | 0.0073144 | 0.015055 | 0.0040825 | 0.0089443 | 0.010328 | 0.01472 | 0.019408 | 0.005164 | 0.016432 | 0.0064936 | 0.040825 | 0.013906 | 0.0041677 | 0.002582 | 0.013784 | 0.014447 | 0.0089443 | 0.010954 | 0.005164 | 0.016021
| Rel.Std.Dev. |  | 2.474% | 2.346% | 0.684% | 2.046% | 1.574% | 1.24% | 1.494% | 3.115% | 0.8303% | 1.825% | 2.094% | 2.786% | 3.947% | 1.047% | 3.459% | 1.296% | 9.798% | 2.665% | 0.8644% | 0.5147% | 2.902% | 3.198% | 1.754% | 2.191% | 1.026% | 3.215%
| PDM3 |  | -0.278% | 0.3935% | 4.431% | 2.557% | -3.636% | 11.37% | -1.386% | -2.628% | -0.9496% | -1.285% | -0.6138% | 6.437% | -0.9496% | -0.6138% | -4.307% | 0.9643% | -16.06% | 5.128% | -2.87% | 1.065% | -4.307% | -8.995% | 2.744% | 0.7292% | 1.401% | 0.3935%
| Z-Score (Absolute) |  | 0 | 0.27 | 1.89 | 1.14 | 1.35 | 4.68 | 0.45 | 0.94 | 0.27 | 0.4 | 0.13 | 2.7 | 0.27 | 0.13 | 1.62 | 0.5 | NA | 2.17 | 1.04 | 0.54 | 1.62 | 3.5 | 1.21 | 0.4 | 0.67 | 0.27
Results from laboratory 1.23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 241b (Certified Value 1.98 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.24 | 2.07 | 2.0858 | 0.2 | 1.9 | 1.93 | < 5 | 2.1 | 2 | 2.06 | 2.04 | 1.99 | < 5 | 2 | 1.8 | 1.87 | 6.14 | 2 | 1.7415 | 1.93 | 1.8488 | 1.95 | 1.97 | 2 | 2.13 |  |  |  |  | 1
| 1 | 2 | 2.26 | 2.13 | 2.1063 | 0.3 | 1.8 | 2.16 | < 5 | 2.1 | 1.96 | 2.02 | 1.94 | 2.08 | < 5 | 2 | 1.8 | 1.9 | 7.79 | 2 | 1.6431 | 1.78 | 1.862 | 1.94 | 2.08 | 2 | 2.05 |  |  |  |  | 32
| 1 | 3 | 2.22 | 2.21 | 2.009 | 0.2 | 1.9 | 1.98 | < 5 | 2.2 | 2.03 | 1.98 | 1.98 | 2 | < 5 | 2 | 1.8 | 1.9 | 8.32 | 2 | 1.701 | 1.97 | 1.8336 | 2.02 | 2.11 | 2 | 2.04 |  |  |  |  | 16
| 1 | 4 | 2.29 | 2.02 | 2.0642 | 0.4 | 1.7 | 2.1 | < 5 | 2.2 | 1.95 | 2 | 1.92 | 1.98 | < 5 | 1.9 | 1.8 | 1.91 | 8.76 | 2.5 | 1.6711 | 1.91 | 1.9041 | 1.86 | 2.01 | 1.8 | 2.06 |  |  |  |  | 1.9844
| 1 | 5 | 2.81 | 2.08 | 2.0178 | 0.6 | 1.9 | 1.94 | < 5 | 2.2 | 1.97 | 2.09 | 1.96 | 2.04 | < 5 | 2 | 1.8 | 1.86 | 8.2 | 2 | 1.7491 | 1.96 | 1.8159 | 2 | 2.14 | 1.9 | 2.15 |  |  |  |  | 53
| 1 | 6 | 2.38 | 2.02 | 2.076 | 0.7 | 1.9 | 2.09 | < 5 | 2.2 | 1.97 | 2.15 | 1.93 | 2.04 | < 5 | 1.9 | 1.8 | 1.88 | 10.78 | 2.5 | 1.7374 | 1.87 | 1.8362 | 1.92 | 2.25 | 1.8 | 2.23
| Mean |  | 2.3667 | 2.0883 | 2.0599 | 0.4 | 1.85 | 2.0333 |  | 2.1667 | 1.98 | 2.05 | 1.9617 | 2.0217 |  | 1.9667 | 1.8 | 1.8867 | 8.3317 | 2.1667 | 1.7072 | 1.9033 | 1.8501 | 1.9483 | 2.0933 | 1.9167 | 2.11
| Median |  | 2.275 | 2.075 | 2.0701 | 0.35 | 1.9 | 2.035 |  | 2.2 | 1.97 | 2.04 | 1.95 | 2.02 |  | 2 | 1.8 | 1.89 | 8.26 | 2 | 1.7192 | 1.92 | 1.8425 | 1.945 | 2.095 | 1.95 | 2.095
| Std Dev. |  | 0.22429 | 0.072503 | 0.038631 | 0.20976 | 0.083666 | 0.095847 |  | 0.05164 | 0.029665 | 0.063246 | 0.044008 | 0.038166 |  | 0.05164 | 0 | 0.019664 | 1.5028 | 0.2582 | 0.043135 | 0.070333 | 0.030641 | 0.057417 | 0.099331 | 0.098319 | 0.074027
| Rel.Std.Dev. |  | 9.477% | 3.472% | 1.875% | 52.44% | 4.522% | 4.714% |  | 2.383% | 1.498% | 3.085% | 2.243% | 1.888% |  | 2.626% | 0% | 1.042% | 18.04% | 11.92% | 2.527% | 3.695% | 1.656% | 2.947% | 4.745% | 5.13% | 3.508%
| PDM3 |  | 19.27% | 5.239% | 3.804% | -79.84% | -6.772% | 2.467% |  | 9.186% | -0.2204% | 3.307% | -1.144% | 1.879% |  | -0.8923% | -9.291% | -4.924% | 319.9% | 9.186% | -13.97% | -4.084% | -6.767% | -1.816% | 5.491% | -3.412% | 6.331%
| Z-Score (Absolute) |  | 1.97 | 0.38 | 0.22 | 9.24 | 0.98 | 0.07 | 2.73 | 0.83 | 0.24 | 0.16 | 0.34 | 0 | 2.73 | 0.31 | 1.26 | 0.77 | 35.96 | 0.83 | 1.79 | 0.67 | 0.98 | 0.42 | 0.41 | 0.6 | 0.5
| Analytical results for Sc in OREAS 241b (Certified Value 35.3 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 37.2 | 35 | 29.09 | 33.552 | 35.2177 | 36 | 47 | 33 | 33.9 | 34 | 35.5 | 37.1 | 36.2 | 35.6 | 34 | 34 | 35 | 33 | 35.8 | 43 | 36.9799 | 36 | 33.1686 | 36.2 | 36.9 | 34.4 | 37.2 |  |  | 1
| 1 | 2 | 37.9 | 34.7 | 29.62 | 35.2608 | 36.3993 | 35 | 45.8 | 33 | 35.5 | 34 | 34.8 | 36.5 | 35.9 | 36.3 | 35.7 | 32 | 34 | 33 | 36.7 | 43 | 36.6419 | 34.2 | 32.9891 | 35.6 | 35.9 | 35 | 35.7 |  |  | 33
| 1 | 3 | 37.9 | 35.5 | 30.16 | 34.5888 | 34.6842 | 35 | 42 | 34 | 35.1 | 34 | 35.3 | 36.6 | 36.1 | 35.6 | 34.7 | 33 | 35 | 33 | 35.5 | 44 | 36.5458 | 36.6 | 33.957 | 36.4 | 36.6 | 35 | 35 |  |  | 16
| 1 | 4 | 38.2 | 34.8 | 29.01 | 34.608 | 35.1524 | 36 | 41.7 | 33 | 35 | 35 | 35.2 | 36.8 | 36.8 | 34.9 | 34.1 | 34 | 34 | 33 | 35.6 | 44 | 35.8259 | 36.1 | 32.8895 | 35 | 35.3 | 34.1 | 36 |  |  | 35.2829
| 1 | 5 | 39.8 | 35.1 | 29.48 | 34.6352 | 35.1407 | 36 | 41.2 | 34 | 35.7 | 33 | 35 | 36.3 | 37.4 | 36.2 | 34.4 | 34 | 34 | 34 | 35.1 | 40 | 36.3941 | 36.5 | 33.6759 | 36.4 | 36.4 | 35 | 37.5 |  |  | 54
| 1 | 6 | 39.1 | 35.2 | 28.79 | 34.992 | 35.9381 | 35 | 38.6 | 34 | 36 | 34 | 35.6 | 36.8 | 38.6 | 35 | 33.5 | 34 | 34 | 33 | 37.3 | 44 | 36.3245 | 36.1 | 32.4501 | 36 | 37.8 | 34.3 | 39.1
| Mean |  | 38.35 | 35.05 | 29.3583 | 34.6061 | 35.4221 | 35.5 | 42.7167 | 33.5 | 35.2 | 34 | 35.2333 | 36.6833 | 36.8333 | 35.6 | 34.4 | 33.5 | 34.3333 | 33.1667 | 36 | 43 | 36.452 | 35.9167 | 33.1884 | 35.9333 | 36.4833 | 34.6333 | 36.75
| Median |  | 38.05 | 35.05 | 29.285 | 34.6216 | 35.1851 | 35.5 | 41.85 | 33.5 | 35.3 | 34 | 35.25 | 36.7 | 36.5 | 35.6 | 34.25 | 34 | 34 | 33 | 35.7 | 43.5 | 36.47 | 36.1 | 33.0789 | 36.1 | 36.5 | 34.7 | 36.6
| Std Dev. |  | 0.93968 | 0.2881 | 0.49845 | 0.58135 | 0.62599 | 0.54772 | 3.1192 | 0.54772 | 0.73756 | 0.63246 | 0.30111 | 0.27869 | 1.025 | 0.5831 | 0.75366 | 0.83666 | 0.5164 | 0.40825 | 0.82946 | 1.5492 | 0.38346 | 0.87502 | 0.54829 | 0.5465 | 0.85654 | 0.41312 | 1.4843
| Rel.Std.Dev. |  | 2.45% | 0.822% | 1.698% | 1.68% | 1.767% | 1.543% | 7.302% | 1.635% | 2.095% | 1.86% | 0.8546% | 0.7597% | 2.783% | 1.638% | 2.191% | 2.497% | 1.504% | 1.231% | 2.304% | 3.603% | 1.052% | 2.436% | 1.652% | 1.521% | 2.348% | 1.193% | 4.039%
| PDM3 |  | 8.693% | -0.66% | -16.79% | -1.918% | 0.3946% | 0.6154% | 21.07% | -5.053% | -0.2348% | -3.636% | -0.1404% | 3.969% | 4.394% | 0.8989% | -2.502% | -5.053% | -2.691% | -5.998% | 2.033% | 21.87% | 3.314% | 1.796% | -5.936% | 1.844% | 3.402% | -1.841% | 4.158%
| Z-Score (Absolute) |  | 1.86 | 0.24 | 3.85 | 0.52 | 0 | 0.05 | 4.63 | 1.22 | 0.14 | 0.9 | 0.12 | 0.8 | 0.9 | 0.11 | 0.65 | 1.22 | 0.69 | 1.43 | 0.37 | 4.81 | 0.65 | 0.31 | 1.42 | 0.32 | 0.67 | 0.5 | 0.84
| Analytical results for Se in OREAS 241b (Indicative Value 0.8 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 2 | 0.6 | < 1.5 | 0.4 | < 1 | < 0.5 | < 5 | 1 | 1 | 1 | 2 | < 5 | < 2 | < 1 | 1 | 8.07 | < 5 | 1 | 0.5923 | 2.9 | < 2 | < 1 | 0.5 | [6 empty] | 1
| 1 | 2 | < 2 | 0.6 | < 1.5 | 0.5 | < 1 | 0.7 | < 5 | 1 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 8.23 | < 5 | 1 | 0.5907 | 3.2 | < 2 | < 1 | 0.5 | [6 empty] | 7
| 1 | 3 | < 2 | 0.6 | < 1.5 | 0.4 | < 1 | 0.6 | < 5 | 2 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 3.31 | < 5 | 1 | 0.5585 | 1.9 | < 2 | < 1 | 0.6 | [6 empty] | 16
| 1 | 4 | < 2 | < 0.5 | < 1.5 | 0.4 | < 1 | 0.5 | < 5 | 2 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 6.49 | < 5 | 1 | 0.6253 | 2 | < 2 | < 1 | < 0.5 | [6 empty] | 0.80073
| 1 | 5 | < 2 | 0.7 | < 1.5 | 0.5 | < 1 | 0.5 | < 5 | 2 | 1 | < 1 | 1 | < 5 | < 2 | < 1 | 1 | 6.16 | < 5 | 1 | 0.6083 | 2.3 | < 2 | < 1 | 0.5 | [6 empty] | 9
| 1 | 6 | < 2 | 0.8 | < 1.5 | 0.4 | < 1 | 0.6 | < 5 | 1 | 1 | 1 | 1 | < 5 | < 2 | < 1 | 1 | 9.76 | < 5 | 1 | 0.5929 | 2.9 | < 2 | < 1 | 0.6
| Mean |  |  | 0.66 |  | 0.43333 |  | 0.58 |  | 1.5 | 1 | 1 | 1.1667 |  |  |  | 1 | 7.0033 |  | 1 | 0.59467 | 2.5333 |  |  | 0.54
| Median |  |  | 0.6 |  | 0.4 |  | 0.6 |  | 1.5 | 1 | 1 | 1 |  |  |  | 1 | 7.28 |  | 1 | 0.5926 | 2.6 |  |  | 0.5
| Std Dev. |  |  | 0.089443 |  | 0.05164 |  | 0.083666 |  | 0.54772 | 0 | 0 | 0.40825 |  |  |  | 0 | 2.2297 |  | 0 | 0.022166 | 0.53914 |  |  | 0.054772
| Rel.Std.Dev. |  |  | 13.55% |  | 11.92% |  | 14.43% |  | 36.51% | 0% | 0% | 34.99% |  |  |  | 0% | 31.84% |  | 0% | 3.728% | 21.28% |  |  | 10.14%
| PDM3 |  |  | -17.57% |  | -45.88% |  | -27.57% |  | 87.33% | 24.89% | 24.89% | 45.7% |  |  |  | 24.89% | 774.6% |  | 24.89% | -25.73% | 216.4% |  |  | -32.56%
| Z-Score (Absolute) |  | 0 | 0.58 | 0.35 | 0.8 | 0.71 | 0.67 | 2.13 | 0.71 | 0 | 0.12 | 0.24 | 2.13 | 0 | 0.71 | 0 | 8.52 | 2.13 | 0 | 0.58 | 2.18 | 0 | 0.71 | 0.72
| Analytical results for Sm in OREAS 241b (Certified Value 2.61 ppm) | [30 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [21 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | [21 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | [21 empty] | 2
3
| 1 | 1 | 2.53 | 2.09 | 2.2956 | 2.6 | 3.1 | 2.78 | 2.9 | 2.5503 | [21 empty] | 1
| 1 | 2 | 2.55 | 2.12 | 2.449 | 2.5 | 3 | 2.71 | 2.9 | 2.6309 | [21 empty] | 18
| 1 | 3 | 2.56 | 2.08 | 2.3665 | 2.5 | 2.7 | 2.66 | 2.7 | 2.6438 | [21 empty] | 16
| 1 | 4 | 2.48 | 2.02 | 2.4175 | 2.4 | 2.6 | 2.65 | 2.7 | 2.6331 | [21 empty] | 2.6059
| 1 | 5 | 2.61 | 2.06 | 2.3923 | 2.7 | 2.7 | 2.69 | 2.7 | 2.5635 | [21 empty] | 55
| 1 | 6 | 2.61 | 2.03 | 2.363 | 2.3 | 2.5 | 2.62 | 2.6 | 2.5941
| Mean |  | 2.5567 | 2.0667 | 2.3806 | 2.5 | 2.7667 | 2.685 | 2.75 | 2.6026
| Median |  | 2.555 | 2.07 | 2.3794 | 2.5 | 2.7 | 2.675 | 2.7 | 2.6125
| Std Dev. |  | 0.049666 | 0.037771 | 0.052743 | 0.14142 | 0.23381 | 0.056125 | 0.12247 | 0.039396
| Rel.Std.Dev. |  | 1.943% | 1.828% | 2.216% | 5.657% | 8.451% | 2.09% | 4.454% | 1.514%
| PDM3 |  | -1.891% | -20.69% | -8.646% | -4.065% | 6.168% | 3.034% | 5.528% | -0.1275%
| Z-Score (Absolute) |  | 0.11 | 2.51 | 0.97 | 0.39 | 0.91 | 0.52 | 0.83 | 0.11
| Analytical results for Sn in OREAS 241b (Certified Value 1.32 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
2
| 1 | 1 | 1.6 | 1.3 | 3.3 | < 2.5 | 1 | 1.67 | 1.6 | 1.2 | < 10 | 1.4 | 1.2 | 1.2 | 1.1 | 1.2 | < 10 | 1.3 | 1.1 | 1.2 | 2.06 | 1.5 | 1.0844 | 1.3 | 0.7865 | 1.3 | 1.3 | 1.3 | 1.9 |  |  | 1
| 1 | 2 | 1.6 | 1.4 | 2.97 | < 2.5 | 1 | 1.69 | 1.6 | 1.4 | < 10 | 1.4 | 1.2 | 1.3 | 1.3 | 1.3 | < 10 | 1.3 | 1.1 | 1.2 | 2.14 | 1.5 | 1.0539 | 1.2 | 0.7769 | 1.3 | 1.4 | 1.3 | 1.9 |  |  | 19
| 1 | 3 | 1.6 | 1.3 | 4.68 | < 2.5 | 1 | 1.46 | 1.6 | 1.3 | < 10 | 1.4 | 1.2 | 1.2 | 1.2 | 1.2 | < 10 | 1.3 | 1.1 | 1.2 | 2.32 | 1.5 | 1.1142 | 1.2 | 0.794 | 1.3 | 1.4 | 1.3 | 1.9 |  |  | 16
| 1 | 4 | 1.6 | 1.3 | 5.23 | < 2.5 | 1 | 1.56 | 1.7 | 1.2 | < 10 | 1.4 | 1.2 | 1.2 | 1.1 | 1.2 | < 10 | 1.1 | 1.1 | 1.2 | 2.1 | 1.5 | 1.073 | 1.3 | 0.8636 | 1.3 | 1.4 | 1.2 | 1.9 |  |  | 1.3151
| 1 | 5 | 1.7 | 1.3 | 3.67 | < 2.5 | 1 | 1.49 | 1.6 | 1.2 | < 10 | 1.3 | 1.2 | 1.3 | 1.2 | 1.2 | < 10 | 1.2 | 1.1 | 1.2 | 2.32 | 1.5 | 1.1138 | 1.3 | 0.7295 | 1.7 | 1.4 | 1.3 | 1.9 |  |  | 56
| 1 | 6 | 1.7 | 1.3 | 2.66 | < 2.5 | 1 | 1.36 | 1.7 | 1.3 | < 10 | 1.4 | 1.2 | 1.3 | 1.1 | 1.2 | < 10 | 1.2 | 1 | 1.3 | 3.69 | 1.5 | 1.107 | 1.2 | 0.7876 | 1.5 | 1.5 | 1.2 | 2
| Mean |  | 1.6333 | 1.3167 | 3.7517 |  | 1 | 1.5383 | 1.6333 | 1.2667 |  | 1.3833 | 1.2 | 1.25 | 1.1667 | 1.2167 |  | 1.2333 | 1.0833 | 1.2167 | 2.4383 | 1.5 | 1.091 | 1.25 | 0.78968 | 1.4 | 1.4 | 1.2667 | 1.9167
| Median |  | 1.6 | 1.3 | 3.485 |  | 1 | 1.525 | 1.6 | 1.25 |  | 1.4 | 1.2 | 1.25 | 1.15 | 1.2 |  | 1.25 | 1.1 | 1.2 | 2.23 | 1.5 | 1.0957 | 1.25 | 0.78705 | 1.3 | 1.4 | 1.3 | 1.9
| Std Dev. |  | 0.05164 | 0.040825 | 1.0061 |  | 0 | 0.12734 | 0.05164 | 0.08165 |  | 0.040825 | 0 | 0.054772 | 0.08165 | 0.040825 |  | 0.08165 | 0.040825 | 0.040825 | 0.6231 | 0 | 0.024727 | 0.054772 | 0.043086 | 0.16733 | 0.063246 | 0.05164 | 0.040825
| Rel.Std.Dev. |  | 3.162% | 3.101% | 26.82% |  | 0% | 8.278% | 3.162% | 6.446% |  | 2.951% | 0% | 4.382% | 6.999% | 3.355% |  | 6.62% | 3.768% | 3.355% | 25.55% | 0% | 2.266% | 4.382% | 5.456% | 11.95% | 4.518% | 4.077% | 2.13%
| PDM3 |  | 24.2% | 0.1226% | 185.3% |  | -23.96% | 16.98% | 24.2% | -3.68% |  | 5.192% | -8.749% | -4.947% | -11.28% | -7.482% |  | -6.214% | -17.62% | -7.482% | 85.42% | 14.06% | -17.04% | -4.947% | -39.95% | 6.459% | 6.459% | -3.68% | 45.75%
| Z-Score (Absolute) |  | 1.97 | 0.14 | 14.22 | 0.24 | NA | 1.43 | 1.97 | 0.14 | 21.43 | 0.53 | 0.53 | 0.24 | 0.72 | 0.43 | 21.43 | 0.34 | 1.2 | 0.43 | 6.63 | 1.2 | 1.16 | 0.24 | 2.9 | 0.63 | 0.63 | 0.14 | 3.61
Results from laboratory 1.07 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 241b (Certified Value 156 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 146.6 | 152.24 | 168 | 150.6554 | 161.3228 | 148 | 163 | 155 | 148.16 | 151 | 154 | 158 | 175.5 | 160 | 162 | 163 | 154.7 | 144 | 160 | 88.87 | 153.7005 | 154 | 157.0726 | 159 | 152.2684 | 154 | 153.5 | 159.3 | 165 | 1
| 1 | 2 | 146.5 | 155.12 | 169 | 158.2112 | 165.6179 | 142 | 162 | 154 | 154 | 153 | 155 | 159 | 169 | 160 | 162.5 | 156 | 152.6 | 141 | 161 | 86.43 | 150.9835 | 152 | 156.0587 | 149.5 | 152.5047 | 152 | 154.2 | 156.1 | 162 | 20
| 1 | 3 | 145 | 153.29 | 165 | 155.6534 | 158.2185 | 146 | 164 | 159 | 153.03 | 151 | 155 | 164.5 | 165 | 158.5 | 164 | 161 | 153.7 | 142 | 156 | 106.73 | 151.6105 | 154 | 159.1296 | 162 | 153.818 | 160 | 154.5 | 160.8 | 155 | 16
| 1 | 4 | 147.4 | 152.23 | 160 | 155.8298 | 160.1999 | 149 | 152 | 157 | 152.4 | 149 | 154 | 163.5 | 170.5 | 154 | 160 | 167 | 150.5 | 145 | 158 | 109.22 | 151.544 | 152 | 151.7911 | 157.5 | 151.9062 | 157 | 156.1 | 159.3 | 160 | 155.6961
| 1 | 5 | 142.9 | 153.07 | 160 | 158.9364 | 160.9939 | 143 | 158 | 161 | 155.03 | 150 | 156 | 161.5 | 178 | 171 | 162 | 164 | 150.9 | 147 | 156 | 96.04 | 149.6725 | 153 | 152.6942 | 161.5 | 148.2952 | 159 | 157.7 | 158.7 | 170 | 57
| 1 | 6 | 141.1 | 150.74 | 156 | 156.9176 | 163.7128 | 151 | 159 | 163 | 153.99 | 148 | 157 | 162 | 172 | 157.5 | 161.5 | 167 | 149.9 | 143 | 162 | 64.59 | 150.195 | 153 | 154.8414 | 153 | 152.1752 | 160 | 160.9 | 159.1 | 182
| Mean |  | 144.9167 | 152.7817 | 163 | 156.034 | 161.6776 | 146.5 | 159.6667 | 158.1667 | 152.7683 | 150.3333 | 155.1667 | 161.4167 | 171.6667 | 160.1667 | 162 | 163 | 152.05 | 143.6667 | 158.8333 | 91.98 | 151.2843 | 153 | 155.2646 | 157.0833 | 151.828 | 157 | 156.15 | 158.8833 | 165.6667
| Median |  | 145.75 | 152.655 | 162.5 | 156.3737 | 161.1583 | 147 | 160.5 | 158 | 153.51 | 150.5 | 155 | 161.75 | 171.25 | 159.25 | 162 | 163.5 | 151.75 | 143.5 | 159 | 92.455 | 151.2638 | 153 | 155.45 | 158.25 | 152.2218 | 158 | 155.3 | 159.2 | 163.5
| Std Dev. |  | 2.4555 | 1.4543 | 5.1381 | 2.9346 | 2.6234 | 3.5071 | 4.4121 | 3.4881 | 2.4317 | 1.7512 | 1.169 | 2.5183 | 4.644 | 5.7504 | 1.3038 | 4.1473 | 1.9181 | 2.1602 | 2.5626 | 16.2601 | 1.4062 | 0.89443 | 2.7455 | 4.9337 | 1.8558 | 3.3466 | 2.7754 | 1.5394 | 9.4375
| Rel.Std.Dev. |  | 1.694% | 0.9519% | 3.152% | 1.881% | 1.623% | 2.394% | 2.763% | 2.205% | 1.592% | 1.165% | 0.7534% | 1.56% | 2.705% | 3.59% | 0.8048% | 2.544% | 1.261% | 1.504% | 1.613% | 17.68% | 0.9295% | 0.5846% | 1.768% | 3.141% | 1.222% | 2.132% | 1.777% | 0.9689% | 5.697%
| PDM3 |  | -6.923% | -1.872% | 4.691% | 0.217% | 3.842% | -5.906% | 2.55% | 1.587% | -1.88% | -3.444% | -0.34% | 3.674% | 10.26% | 2.871% | 4.049% | 4.691% | -2.342% | -7.726% | 2.015% | -40.92% | -2.834% | -1.732% | -0.2771% | 0.891% | -2.484% | 0.8375% | 0.2916% | 2.047% | 6.404%
| Z-Score (Absolute) |  | 1.85 | 0.55 | 1.13 | 0.02 | 0.91 | 1.59 | 0.58 | 0.33 | 0.56 | 0.96 | 0.16 | 0.87 | 2.55 | 0.66 | 0.96 | 1.13 | 0.67 | 2.05 | 0.44 | 10.55 | 0.8 | 0.52 | 0.15 | 0.15 | 0.71 | 0.14 | 0 | 0.45 | 1.57
| Analytical results for Ta in OREAS 241b (Certified Value 0.25 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.2 | 0.25 | 0.41447 | < 0.1 | 0.27 | 0.36 | 0.23 | 12 | 0.4 | 0.23 | 0.26 | 0.22 | 0.27 | 0.22 | 0.2 | 0.25 | < 0.5 | 0.2122 | 0.25 | 0.5091 | 0.42 | 0.67 | 0.26 | 0.26 | [5 empty] | 1
| 1 | 2 | 0.15 | 0.26 | 0.35551 | < 0.1 | 0.38 | 0.4 | 0.25 | 11 | 0.4 | 0.23 | 0.26 | 0.24 | 0.28 | 0.24 | 0.2 | 0.24 | < 0.5 | 0.21684 | 0.23 | 0.3951 | 0.42 | 0.61 | 0.26 | 0.25 | [5 empty] | 21
| 1 | 3 | 0.17 | 0.26 | 0.33573 | < 0.1 | 0.25 | 0.41 | 0.23 | 12 | 0.4 | 0.24 | 0.25 | 0.23 | 0.29 | 0.23 | 0.2 | 0.24 | < 0.5 | 0.21915 | 0.26 | 0.4256 | 0.42 | 0.66 | 0.26 | 0.24 | [5 empty] | 16
| 1 | 4 | 0.1 | 0.26 | 0.32656 | < 0.1 | 0.25 | 0.39 | 0.24 | 12 | 0.4 | 0.24 | 0.26 | 0.23 | 0.29 | 0.21 | 0.2 | 0.24 | < 0.5 | 0.22044 | 0.25 | 0.4393 | 0.37 | 0.71 | 0.26 | 0.25 | [5 empty] | 0.25106
| 1 | 5 | 0.74 | 0.26 | 0.32176 | < 0.1 | 0.24 | 0.41 | 0.24 | 11 | 0.4 | 0.24 | 0.27 | 0.24 | 0.28 | 0.23 | 0.2 | 0.24 | < 0.5 | 0.21124 | 0.26 | 0.3674 | 0.61 | 0.68 | 0.27 | 0.25 | [5 empty] | 58
| 1 | 6 | 0.24 | 0.25 | 0.32468 | < 0.1 | 0.22 | 0.37 | 0.25 | 11 | 0.4 | 0.23 | 0.26 | 0.23 | 0.25 | 0.21 | 0.2 | 0.24 | < 0.5 | 0.2174 | 0.24 | 0.3594 | 0.46 | 0.71 | 0.26 | 0.27
| Mean |  | 0.26667 | 0.25667 | 0.34645 |  | 0.26833 | 0.39 | 0.24 | 11.5 | 0.4 | 0.235 | 0.26 | 0.23167 | 0.27667 | 0.22333 | 0.2 | 0.24167 |  | 0.21621 | 0.24833 | 0.41598 | 0.45 | 0.67333 | 0.26167 | 0.25333
| Median |  | 0.185 | 0.26 | 0.33115 |  | 0.25 | 0.395 | 0.24 | 11.5 | 0.4 | 0.235 | 0.26 | 0.23 | 0.28 | 0.225 | 0.2 | 0.24 |  | 0.21712 | 0.25 | 0.41035 | 0.42 | 0.675 | 0.26 | 0.25
| Std Dev. |  | 0.23662 | 0.005164 | 0.035505 |  | 0.057067 | 0.020976 | 0.0089443 | 0.54772 | 6.0809e-17 | 0.0054772 | 0.0063246 | 0.0075277 | 0.015055 | 0.012111 | 3.0405e-17 | 0.0040825 |  | 0.0037183 | 0.01169 | 0.055326 | 0.083427 | 0.037238 | 0.0040825 | 0.010328
| Rel.Std.Dev. |  | 88.73% | 2.012% | 10.25% |  | 21.27% | 5.379% | 3.727% | 4.763% | 1.52e-14% | 2.331% | 2.433% | 3.249% | 5.442% | 5.423% | 1.52e-14% | 1.689% |  | 1.72% | 4.708% | 13.3% | 18.54% | 5.53% | 1.56% | 4.077%
| PDM3 |  | 6.215% | 2.232% | 38% |  | 6.879% | 55.34% | -4.406% | 4481% | 59.32% | -6.398% | 3.56% | -7.726% | 10.2% | -11.04% | -20.34% | -3.743% |  | -13.88% | -1.087% | 65.69% | 79.24% | 168.2% | 4.224% | 0.9043%
| Z-Score (Absolute) |  | 0.34 | 0 | 3.03 | 6.97 | 0.39 | 4.5 | 0.56 | NA | NA | 0.73 | 0.11 | 0.84 | 0.67 | 1.12 | NA | 0.51 | 0.22 | 1.36 | 0.28 | 5.37 | 6.52 | 14.05 | 0.17 | 0.11
Results from laboratories 1.13 and 1.23 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.12 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 241b (Certified Value 0.53 ppm) | [30 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 0.61 | 0.55 | 0.5 | 0.4857 | 0.6 | 0.61 | 0.53 | 0.55 | 3.1273 | 0.54 | [19 empty] | 1
| 1 | 2 | 0.61 | 0.58 | 0.51 | 0.49685 | 0.6 | 0.6 | 0.54 | 0.55 | 3.3006 | 0.53 | [19 empty] | 22
| 1 | 3 | 0.63 | 0.56 | 0.5 | 0.50204 | 0.6 | 0.55 | 0.54 | 0.5 | 3.3376 | 0.54 | [19 empty] | 16
| 1 | 4 | 0.62 | 0.55 | 0.49 | 0.48751 | 0.6 | 0.54 | 0.54 | 0.55 | 3.3481 | 0.53 | [19 empty] | 0.53307
| 1 | 5 | 0.65 | 0.56 | 0.51 | 0.49859 | 0.6 | 0.55 | 0.53 | 0.5 | 3.3161 | 0.56 | [19 empty] | 59
| 1 | 6 | 0.64 | 0.57 | 0.5 | 0.48841 | 0.6 | 0.52 | 0.53 | 0.55 | 3.2707 | 0.57
| Mean |  | 0.62667 | 0.56167 | 0.50167 | 0.49318 | 0.6 | 0.56167 | 0.535 | 0.53333 | 3.2834 | 0.545
| Median |  | 0.625 | 0.56 | 0.5 | 0.49263 | 0.6 | 0.55 | 0.535 | 0.55 | 3.3083 | 0.54
| Std Dev. |  | 0.01633 | 0.01169 | 0.0075277 | 0.0068136 | 0 | 0.035449 | 0.0054772 | 0.02582 | 0.081263 | 0.016432
| Rel.Std.Dev. |  | 2.606% | 2.081% | 1.501% | 1.382% | 0% | 6.311% | 1.024% | 4.841% | 2.475% | 3.015%
| PDM3 |  | 17.56% | 5.364% | -5.892% | -7.483% | 12.55% | 5.364% | 0.3613% | 0.04869% | 515.9% | 2.237%
| Z-Score (Absolute) |  | 3.3 | 0.67 | 1.75 | 2.1 | NA | 0.67 | 0.4 | 0.47 | 110.79 | 0
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 241b (Certified Value 0.13 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.2 | < 0.2 | 0.20738 | < 0.1 | < 0.1 | < 0.2 | 14 | < 0.1 | 0.16 | 0.14 | 0.11 | 0.21 | < 5 | 0.1 | < 0.5 | 0.13 | < 1 | 0.14108 | 0.13 | 0.0863 | 0.06 | 0.12 | 0.15 | 0.11 | [5 empty] | 1
| 1 | 2 | 0.11 | < 0.2 | 0.20234 | < 0.1 | 0.2 | < 0.2 | 15 | 0.1 | 0.17 | 0.12 | 0.11 | 0.14 | < 5 | 0.1 | < 0.5 | 0.12 | < 1 | 0.11977 | 0.14 | 0.0945 | 0.14 | 0.16 | 0.16 | 0.08 | [5 empty] | 23
| 1 | 3 | 0.14 | < 0.2 | 0.20972 | < 0.1 | < 0.1 | < 0.2 | 14 | < 0.1 | 0.16 | 0.11 | 0.13 | 0.15 | < 5 | 0.1 | < 0.5 | 0.12 | < 1 | 0.11167 | 0.14 | 0.1356 | 0.17 | 0.15 | 0.14 | 0.12 | [5 empty] | 16
| 1 | 4 | 0.17 | < 0.2 | 0.19397 | < 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.13 | 0.15 | 0.14 | 0.12 | < 5 | 0.1 | < 0.5 | 0.1 | < 1 | 0.12627 | 0.14 | 0.0863 | 0.08 | 0.18 | 0.15 | 0.1 | [5 empty] | 0.13309
| 1 | 5 | 0.24 | < 0.2 | 0.17729 | < 0.1 | < 0.1 | < 0.2 | 15 | < 0.1 | 0.14 | 0.14 | 0.11 | 0.21 | < 5 | 0.1 | < 0.5 | 0.12 | < 1 | 0.12096 | 0.14 | 0.1356 | 0.04 | 0.13 | 0.15 | 0.14 | [5 empty] | 60
| 1 | 6 | 0.14 | < 0.2 | 0.18661 | < 0.1 | 0.2 | < 0.2 | 15 | 0.1 | 0.14 | 0.14 | 0.12 | 0.19 | < 5 | 0.1 | < 0.5 | 0.11 | < 1 | 0.14071 | 0.15 | 0.0904 | 0.07 | 0.16 | 0.16 | 0.18
| Mean |  | 0.16667 |  | 0.19622 |  | 0.2 |  | 14.6667 | 0.1 | 0.15 | 0.13333 | 0.12 | 0.17 |  | 0.1 |  | 0.11667 |  | 0.12674 | 0.14 | 0.10478 | 0.093333 | 0.15 | 0.15167 | 0.12167
| Median |  | 0.155 |  | 0.19815 |  | 0.2 |  | 15 | 0.1 | 0.15 | 0.14 | 0.115 | 0.17 |  | 0.1 |  | 0.12 |  | 0.12362 | 0.14 | 0.09245 | 0.075 | 0.155 | 0.15 | 0.115
| Std Dev. |  | 0.047188 |  | 0.012646 |  | 0 |  | 0.5164 | 0 | 0.015492 | 0.015055 | 0.012649 | 0.038471 |  | 1.5202e-17 |  | 0.010328 |  | 0.011917 | 0.0063246 | 0.024063 | 0.050465 | 0.021909 | 0.0075277 | 0.034881
| Rel.Std.Dev. |  | 28.31% |  | 6.445% |  | 0% |  | 3.521% | 0% | 10.33% | 11.29% | 10.54% | 22.63% |  | 1.52e-14% |  | 8.853% |  | 9.403% | 4.518% | 22.96% | 54.07% | 14.61% | 4.963% | 28.67%
| PDM3 |  | 25.23% |  | 47.43% |  | 50.27% |  | 10920% | -24.86% | 12.7% | 0.1815% | -9.837% | 27.73% |  | -24.86% |  | -12.34% |  | -4.77% | 5.191% | -21.27% | -29.87% | 12.7% | 13.96% | -8.584%
| Z-Score (Absolute) |  | 0.65 | 0.8 | 1.3 | 1.89 | NA | 0.8 | NA | NA | 0.29 | 0.07 | 0.36 | 0.73 | 51.5 | NA | 2.47 | 0.44 | 7.92 | 0.22 | 0.07 | 0.69 | 0.94 | 0.29 | 0.33 | 0.33
| Analytical results for Th in OREAS 241b (Certified Value 1.3 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 2.6 | 1.29 | 1.21 | 1.0062 | 1.2906 | 1.3 | 1.67 | 1.39 | 1.25 | 1.3 | 1.22 | 1.27 | 1.24 | 1.23 | 1.27 | 1.1 | 1.4 | 1.6 | 1.3306 | 1.28 | 1.5326 | 1.4 | 1.3 | 1.3 | 1.4 |  |  |  |  | 1
| 1 | 2 | 2.6 | 1.29 | 1.2 | 1.0769 | 1.2746 | 1.3 | 1.64 | 1.4 | 1.31 | 1.3 | 1.23 | 1.32 | 1.26 | 1.24 | 1.34 | 1.1 | 1.36 | 1.5 | 1.3731 | 1.27 | 1.5303 | 1.4 | 1.2 | 1.3 | 1.4 |  |  |  |  | 24
| 1 | 3 | 2.5 | 1.32 | 1.35 | 1.0371 | 1.2751 | 1.3 | 1.49 | 1.41 | 1.31 | 1.3 | 1.21 | 1.31 | 1.24 | 1.2 | 1.34 | 1.1 | 1.37 | 1.5 | 1.394 | 1.31 | 1.533 | 1.5 | 1.3 | 1.3 | 1.3 |  |  |  |  | 16
| 1 | 4 | 1.3 | 1.28 | 1.16 | 1.008 | 1.2626 | 1.3 | 1.5 | 1.4 | 1.27 | 1.3 | 1.23 | 1.29 | 1.2 | 1.2 | 1.35 | 1.1 | 1.38 | 1.6 | 1.3545 | 1.29 | 1.5522 | 1.4 | 1.3 | 1.2 | 1.3 |  |  |  |  | 1.3048
| 1 | 5 | 2.5 | 1.27 | 1.18 | 1.0474 | 1.262 | 1.3 | 1.51 | 1.41 | 1.26 | 1.3 | 1.23 | 1.29 | 1.28 | 1.23 | 1.32 | 1.2 | 1.35 | 1.5 | 1.3922 | 1.31 | 1.51 | 1.5 | 1.3 | 1.2 | 1.4 |  |  |  |  | 61
| 1 | 6 | 2.3 | 1.28 | 1.16 | 1.0426 | 1.3024 | 1.4 | 1.42 | 1.42 | 1.31 | 1.3 | 1.21 | 1.4 | 1.28 | 1.2 | 1.27 | 1.1 | 1.35 | 1.6 | 1.4075 | 1.28 | 1.5252 | 1.8 | 1.4 | 1.2 | 1.5
| Mean |  | 2.3 | 1.2883 | 1.21 | 1.0364 | 1.2779 | 1.3167 | 1.5383 | 1.405 | 1.285 | 1.3 | 1.2217 | 1.3133 | 1.25 | 1.2167 | 1.315 | 1.1167 | 1.3683 | 1.55 | 1.3753 | 1.29 | 1.5305 | 1.5 | 1.3 | 1.25 | 1.3833
| Median |  | 2.5 | 1.285 | 1.19 | 1.0398 | 1.2749 | 1.3 | 1.505 | 1.405 | 1.29 | 1.3 | 1.225 | 1.3 | 1.25 | 1.215 | 1.33 | 1.1 | 1.365 | 1.55 | 1.3826 | 1.285 | 1.5314 | 1.45 | 1.3 | 1.25 | 1.4
| Std Dev. |  | 0.502 | 0.017224 | 0.071554 | 0.02653 | 0.015911 | 0.040825 | 0.096212 | 0.010488 | 0.028107 | 0 | 0.0098319 | 0.045898 | 0.030332 | 0.018619 | 0.036194 | 0.040825 | 0.019408 | 0.054772 | 0.028659 | 0.016733 | 0.013638 | 0.15492 | 0.063246 | 0.054772 | 0.075277
| Rel.Std.Dev. |  | 21.83% | 1.337% | 5.914% | 2.56% | 1.245% | 3.101% | 6.254% | 0.7465% | 2.187% | 0% | 0.8048% | 3.495% | 2.427% | 1.53% | 2.752% | 3.656% | 1.418% | 3.534% | 2.084% | 1.297% | 0.891% | 10.33% | 4.865% | 4.382% | 5.442%
| PDM3 |  | 76.28% | -1.26% | -7.264% | -20.57% | -2.06% | 0.9111% | 17.9% | 7.681% | -1.516% | -0.3663% | -6.37% | 0.6556% | -4.198% | -6.753% | 0.7834% | -14.42% | 4.871% | 18.79% | 5.406% | -1.133% | 17.3% | 14.96% | -0.3663% | -4.198% | 6.021%
| Z-Score (Absolute) |  | NA | 0.01 | 0.76 | 2.44 | 0.11 | NA | 2.4 | 1.12 | 0.04 | NA | 0.65 | 0.23 | 0.38 | 0.7 | 0.25 | NA | 0.76 | NA | 0.83 | 0.01 | 2.33 | NA | NA | NA | NA
Results from laboratories 1.01, 1.07, 1.13, 1.23, 1.27, 1.31, 1.32, 1.33 and 1.34 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 241b (Certified Value 0.592 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | wt.%
|  |  | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.58 | 0.579 | 0.58244 | 0.60661 | 0.343 | 0.6474 | 0.569 | 0.5686 | 0.57 | 0.61 | 0.593 | 0.637 | 0.589 | 0.602 | 0.56 | 0.6164 | 0.575 | 0.593 | 0.675 | 0.58399 | 0.585 | 0.58641 | 0.583 | 0.50499 | 0.57 | 0.6 | 0.59 | 0.65 |  | 1
| 1 | 2 | 0.58 | 0.5981 | 0.58637 | 0.62756 | 0.428 | 0.6414 | 0.58 | 0.5965 | 0.57 | 0.61 | 0.593 | 0.606 | 0.582 | 0.581 | 0.52 | 0.6135 | 0.565 | 0.609 | 0.671 | 0.57035 | 0.583 | 0.57659 | 0.561 | 0.50659 | 0.56 | 0.59 | 0.59 | 0.64 |  | 25
| 1 | 3 | 0.6 | 0.6023 | 0.58747 | 0.60128 | 0.291 | 0.6534 | 0.594 | 0.5838 | 0.57 | 0.607 | 0.624 | 0.6 | 0.576 | 0.59 | 0.54 | 0.6118 | 0.571 | 0.603 | 0.685 | 0.57675 | 0.587 | 0.58363 | 0.594 | 0.52589 | 0.59 | 0.59 | 0.6 | 0.62 |  | 16
| 1 | 4 | 0.6 | 0.5829 | 0.58723 | 0.60787 | 0.341 | 0.6055 | 0.583 | 0.5848 | 0.57 | 0.603 | 0.618 | 0.611 | 0.573 | 0.588 | 0.56 | 0.607 | 0.578 | 0.594 | 0.679 | 0.57702 | 0.578 | 0.57074 | 0.589 | 0.51269 | 0.58 | 0.58 | 0.59 | 0.64 |  | 0.59234
| 1 | 5 | 0.59 | 0.5939 | 0.58802 | 0.60764 | 0.482 | 0.6235 | 0.592 | 0.5963 | 0.57 | 0.603 | 0.599 | 0.629 | 0.637 | 0.583 | 0.55 | 0.6051 | 0.589 | 0.603 | 0.702 | 0.56853 | 0.582 | 0.57493 | 0.604 | 0.50669 | 0.58 | 0.6 | 0.59 | 0.63 |  | 62
| 1 | 6 | 0.59 | 0.578 | 0.58741 | 0.61863 | 0.455 | 0.6355 | 0.599 | 0.5943 | 0.56 | 0.611 | 0.587 | 0.611 | 0.579 | 0.598 | 0.56 | 0.6125 | 0.57 | 0.617 | 0.697 | 0.56808 | 0.579 | 0.5757 | 0.571 | 0.50579 | 0.59 | 0.61 | 0.59 | 0.64
| Mean |  | 0.59 | 0.58903 | 0.58649 | 0.6116 | 0.39 | 0.63445 | 0.58617 | 0.58738 | 0.56833 | 0.60733 | 0.60233 | 0.61567 | 0.58933 | 0.59033 | 0.54833 | 0.61105 | 0.57467 | 0.60317 | 0.68483 | 0.57412 | 0.58233 | 0.578 | 0.58367 | 0.51044 | 0.57833 | 0.595 | 0.59167 | 0.63667
| Median |  | 0.59 | 0.5884 | 0.58732 | 0.60775 | 0.3855 | 0.63845 | 0.5875 | 0.58955 | 0.57 | 0.6085 | 0.596 | 0.611 | 0.5805 | 0.589 | 0.555 | 0.61215 | 0.573 | 0.603 | 0.682 | 0.57355 | 0.5825 | 0.57614 | 0.586 | 0.50664 | 0.58 | 0.595 | 0.59 | 0.64
| Std Dev. |  | 0.0089443 | 0.010411 | 0.0020524 | 0.0096441 | 0.075557 | 0.017517 | 0.01098 | 0.010796 | 0.0040825 | 0.0036148 | 0.015069 | 0.01425 | 0.023989 | 0.0082624 | 0.016021 | 0.0042213 | 0.0083106 | 0.0090866 | 0.012368 | 0.006241 | 0.0034448 | 0.0058632 | 0.015642 | 0.0080496 | 0.01169 | 0.010488 | 0.0040825 | 0.010328
| Rel.Std.Dev. |  | 1.516% | 1.767% | 0.3499% | 1.577% | 19.37% | 2.761% | 1.873% | 1.838% | 0.7183% | 0.5952% | 2.502% | 2.315% | 4.071% | 1.4% | 2.922% | 0.6908% | 1.446% | 1.506% | 1.806% | 1.087% | 0.5916% | 1.014% | 2.68% | 1.577% | 2.021% | 1.763% | 0.69% | 1.622%
| PDM3 |  | -0.3944% | -0.5576% | -0.9871% | 3.252% | -34.16% | 7.11% | -1.042% | -0.8361% | -4.052% | 2.532% | 1.688% | 3.939% | -0.5069% | -0.3381% | -7.429% | 3.159% | -2.983% | 1.828% | 15.62% | -3.075% | -1.689% | -2.42% | -1.464% | -13.83% | -2.364% | 0.4497% | -0.113% | 7.484%
| Z-Score (Absolute) |  | 0.04 | 0.01 | 0.13 | 1.11 | 9.9 | 2.25 | 0.15 | 0.09 | 1.04 | 0.9 | 0.65 | 1.32 | 0.01 | 0.06 | 2.03 | 1.09 | 0.72 | 0.7 | 4.75 | 0.75 | 0.34 | 0.56 | 0.27 | 3.91 | 0.54 | 0.29 | 0.12 | 2.36
| Analytical results for Tl in OREAS 241b (Certified Value 0.31 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.3 | 0.33 | 0.35205 | 0.3 | 0.32 | 0.31 | 0.3 | 0.29 | 0.32 | 0.3 | 0.3 | < 5 | 0.33 | < 0.5 | 0.32 | 0.6 | 0.3 | 0.32318 | 0.31 | 0.312 | 0.27 | 0.3 | 0.3 | 0.33 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.32 | 0.32758 | 0.31 | 0.3 | 0.33 | 0.3 | 0.3 | 0.33 | 0.32 | 0.28 | < 5 | 0.34 | < 0.5 | 0.32 | < 0.1 | 0.3 | 0.32356 | 0.32 | 0.3171 | 0.27 | 0.29 | 0.29 | 0.32 | [5 empty] | 26
| 1 | 3 | 0.3 | 0.32 | 0.31981 | 0.31 | 0.3 | 0.33 | 0.3 | 0.29 | 0.31 | 0.3 | 0.29 | < 5 | 0.34 | < 0.5 | 0.31 | 0.65 | 0.35 | 0.33924 | 0.3 | 0.3144 | 0.27 | 0.29 | 0.29 | 0.31 | [5 empty] | 16
| 1 | 4 | 0.31 | 0.31 | 0.31472 | 0.3 | 0.3 | 0.32 | 0.3 | 0.31 | 0.32 | 0.29 | 0.3 | < 5 | 0.33 | < 0.5 | 0.31 | 0.77 | 0.3 | 0.32258 | 0.3 | 0.3164 | 0.26 | 0.3 | 0.28 | 0.32 | [5 empty] | 0.30994
| 1 | 5 | 0.35 | 0.31 | 0.31131 | 0.3 | 0.34 | 0.33 | 0.3 | 0.3 | 0.33 | 0.32 | 0.29 | < 5 | 0.33 | < 0.5 | 0.31 | 0.73 | 0.3 | 0.33125 | 0.33 | 0.3145 | 0.29 | 0.3 | 0.28 | 0.33 | [5 empty] | 63
| 1 | 6 | 0.31 | 0.33 | 0.32151 | 0.31 | 0.34 | 0.32 | 0.3 | 0.31 | 0.33 | 0.29 | 0.3 | < 5 | 0.34 | < 0.5 | 0.32 | < 0.1 | 0.3 | 0.3273 | 0.32 | 0.3139 | 0.29 | 0.32 | 0.27 | 0.34
| Mean |  | 0.31167 | 0.32 | 0.3245 | 0.305 | 0.31667 | 0.32333 | 0.3 | 0.3 | 0.32333 | 0.30333 | 0.29333 |  | 0.335 |  | 0.315 | 0.6875 | 0.30833 | 0.32785 | 0.31333 | 0.31472 | 0.275 | 0.3 | 0.285 | 0.325
| Median |  | 0.305 | 0.32 | 0.32066 | 0.305 | 0.31 | 0.325 | 0.3 | 0.3 | 0.325 | 0.3 | 0.295 |  | 0.335 |  | 0.315 | 0.69 | 0.3 | 0.32543 | 0.315 | 0.31445 | 0.27 | 0.3 | 0.285 | 0.325
| Std Dev. |  | 0.019408 | 0.0089443 | 0.014622 | 0.0054772 | 0.019664 | 0.008165 | 0 | 0.0089443 | 0.008165 | 0.013663 | 0.008165 |  | 0.0054772 |  | 0.0054772 | 0.076757 | 0.020412 | 0.0064743 | 0.012111 | 0.001828 | 0.012247 | 0.010954 | 0.010488 | 0.010488
| Rel.Std.Dev. |  | 6.227% | 2.795% | 4.506% | 1.796% | 6.21% | 2.525% | 0% | 2.981% | 2.525% | 4.504% | 2.784% |  | 1.635% |  | 1.739% | 11.16% | 6.62% | 1.975% | 3.865% | 0.5808% | 4.454% | 3.651% | 3.68% | 3.227%
| PDM3 |  | 0.5557% | 3.244% | 4.695% | -1.595% | 2.169% | 4.32% | -3.208% | -3.208% | 4.32% | -2.133% | -5.359% |  | 8.084% |  | 1.631% | 121.8% | -0.5198% | 5.778% | 1.093% | 1.54% | -11.27% | -3.208% | -8.048% | 4.858%
| Z-Score (Absolute) |  | 0.2 | 0.35 | 0.64 | 0.64 | 0.13 | 0.57 | NA | 0.97 | 0.57 | 0.75 | 1.4 | 143.3 | 1.33 | 4.24 | 0.02 | 10.51 | 0.42 | 0.86 | 0.09 | 0 | 2.6 | 0.97 | 1.95 | 0.67
Results from laboratory 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 241b (Certified Value 0.32 ppm) | [30 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [22 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.24 | 1.27 | 1.28 | [22 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | [22 empty] | 2
3
| 1 | 1 | 0.32 | 0.29 | 0.2787 | 0.3 | 0.32 | 0.35 | 0.33987 | [22 empty] | 1
| 1 | 2 | 0.32 | 0.29 | 0.28622 | 0.3 | 0.3 | 0.35 | 0.34834 | [22 empty] | 27
| 1 | 3 | 0.33 | 0.29 | 0.28251 | 0.3 | 0.3 | 0.35 | 0.35662 | [22 empty] | 16
| 1 | 4 | 0.33 | 0.29 | 0.27815 | 0.3 | 0.3 | 0.35 | 0.3603 | [22 empty] | 0.31763
| 1 | 5 | 0.35 | 0.29 | 0.28668 | 0.3 | 0.31 | 0.35 | 0.34407 | [22 empty] | 64
| 1 | 6 | 0.33 | 0.29 | 0.27788 | 0.3 | 0.3 | 0.4 | 0.34529
| Mean |  | 0.33 | 0.29 | 0.28169 | 0.3 | 0.305 | 0.35833 | 0.34908
| Median |  | 0.33 | 0.29 | 0.28061 | 0.3 | 0.3 | 0.35 | 0.34681
| Std Dev. |  | 0.010954 | 0 | 0.0040518 | 0 | 0.0083666 | 0.020412 | 0.0078432
| Rel.Std.Dev. |  | 3.32% | 0% | 1.438% | 0% | 2.743% | 5.696% | 2.247%
| PDM3 |  | 3.895% | -8.698% | -11.31% | -5.55% | -3.976% | 12.82% | 9.902%
| Z-Score (Absolute) |  | 0.29 | 0.63 | 0.82 | NA | 0.29 | 0.93 | 0.72
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 241b (Certified Value 0.4 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 0.5 | 0.4 | 0.37 | 0.30818 | 0.41208 | 0.4 | 0.52 | 0.4 | 0.39 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.46 | 0.3 | 0.41 | 0.5 | 0.37562 | 0.4 | 0.3904 | 0.379 | 0.4 | 0.4 | 0.401 |  |  |  | 1
| 1 | 2 | 0.5 | 0.4 | 0.38 | 0.3065 | 0.41788 | 0.4 | 0.52 | 0.38 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.42 | 0.3 | 0.41 | 0.5 | 0.38394 | 0.4 | 0.4031 | 0.379 | 0.4 | 0.4 | 0.402 |  |  |  | 28
| 1 | 3 | 0.6 | 0.4 | 0.39 | 0.29683 | 0.38916 | 0.4 | 0.51 | 0.39 | 0.41 | 0.4 | 0.4 | 0.4 | 0.4 | 0.3 | < 10 | 0.37 | 0.3 | 0.41 | 0.5 | 0.39279 | 0.4 | 0.394 | 0.379 | 0.4 | 0.4 | 0.39 |  |  |  | 16
| 1 | 4 | 0.5 | 0.39 | 0.37 | 0.29624 | 0.38549 | 0.4 | 0.53 | 0.38 | 0.39 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.44 | 0.3 | 0.4 | 0.5 | 0.38412 | 0.4 | 0.402 | 0.377 | 0.4 | 0.3 | 0.4 |  |  |  | 0.39681
| 1 | 5 | 0.6 | 0.41 | 0.38 | 0.31049 | 0.38729 | 0.4 | 0.48 | 0.4 | 0.39 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.37 | 0.3 | 0.4 | 0.4 | 0.38664 | 0.4 | 0.3959 | 0.383 | 0.4 | 0.4 | 0.418 |  |  |  | 65
| 1 | 6 | 0.5 | 0.38 | 0.37 | 0.30419 | 0.40255 | 0.4 | 0.45 | 0.37 | 0.39 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 | < 10 | 0.36 | 0.3 | 0.4 | 0.5 | 0.38943 | 0.4 | 0.3976 | 0.375 | 0.4 | 0.3 | 0.431
| Mean |  | 0.53333 | 0.39667 | 0.37667 | 0.30374 | 0.39907 | 0.4 | 0.50167 | 0.38667 | 0.395 | 0.4 | 0.4 | 0.4 | 0.4 | 0.38333 |  | 0.40333 | 0.3 | 0.405 | 0.48333 | 0.38542 | 0.4 | 0.39717 | 0.37867 | 0.4 | 0.36667 | 0.407
| Median |  | 0.5 | 0.4 | 0.375 | 0.30534 | 0.39586 | 0.4 | 0.515 | 0.385 | 0.39 | 0.4 | 0.4 | 0.4 | 0.4 | 0.4 |  | 0.395 | 0.3 | 0.405 | 0.5 | 0.38538 | 0.4 | 0.39675 | 0.379 | 0.4 | 0.4 | 0.4015
| Std Dev. |  | 0.05164 | 0.010328 | 0.008165 | 0.0059505 | 0.013831 | 6.0809e-17 | 0.030605 | 0.012111 | 0.0083666 | 6.0809e-17 | 6.0809e-17 | 6.0809e-17 | 6.0809e-17 | 0.040825 |  | 0.042269 | 0 | 0.0054772 | 0.040825 | 0.0058618 | 6.0809e-17 | 0.0048202 | 0.0026583 | 6.0809e-17 | 0.05164 | 0.014805
| Rel.Std.Dev. |  | 9.682% | 2.604% | 2.168% | 1.959% | 3.466% | 1.52e-14% | 6.101% | 3.132% | 2.118% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 10.65% |  | 10.48% | 0% | 1.352% | 8.447% | 1.521% | 1.52e-14% | 1.214% | 0.702% | 1.52e-14% | 14.08% | 3.638%
| PDM3 |  | 34.4% | -0.03731% | -5.077% | -23.46% | 0.5696% | 0.8027% | 26.42% | -2.557% | -0.4573% | 0.8027% | 0.8027% | 0.8027% | 0.8027% | -3.397% |  | 1.643% | -24.4% | 2.063% | 21.8% | -2.871% | 0.8027% | 0.08869% | -4.573% | 0.8027% | -7.598% | 2.567%
| Z-Score (Absolute) |  | 14.98 | 0.37 | 2.62 | 10.82 | 0.1 | 0 | 11.43 | 1.5 | 0.56 | 0 | 0 | 0 | 0 | 1.87 | 516.97 | 0.37 | 11.24 | 0.56 | 9.37 | 1.64 | 0 | 0.32 | 2.4 | 0 | 3.75 | 0.79
| Analytical results for V in OREAS 241b (Certified Value 250 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 255 | 246 | 224.86 | 239.7646 | 249.1093 | 165 | 313 | 240 | 242 | 241 | 257 | 249 | 271 | 249 | 249 | 255 | 264 | 249 | 263 | 209.44 | 260.0748 | 250 | 253.8854 | 252 | 232.7654 | 254 | 246 | 264 | 250 | 1
| 1 | 2 | 257 | 247 | 227.24 | 242.1418 | 260.6269 | 192 | 315 | 246 | 256 | 240 | 263 | 246 | 255 | 249 | 243 | 239 | 256 | 243 | 266 | 205.01 | 252.5676 | 245 | 253.6355 | 242 | 231.8516 | 252 | 244 | 265 | 240 | 29
| 1 | 3 | 264 | 252 | 224.57 | 247.4955 | 248.1172 | 142 | 316 | 249 | 241 | 240 | 259 | 258 | 250 | 247 | 250 | 248 | 257 | 246 | 259 | 208.34 | 255.5024 | 250 | 253.6629 | 251 | 236.872 | 257 | 245 | 268 | 239 | 16
| 1 | 4 | 263 | 253 | 223.15 | 241.8501 | 249.8698 | 146 | 298 | 235 | 235 | 242 | 260 | 259 | 260 | 244 | 248 | 260 | 244 | 249 | 258 | 200.49 | 256.5328 | 250 | 252.8627 | 250 | 230.8552 | 248 | 243 | 265 | 218 | 250.4634
| 1 | 5 | 259 | 250 | 224.68 | 247.9327 | 248.6514 | 188 | 308 | 240 | 242 | 241 | 254 | 248 | 267 | 274 | 249 | 255 | 269 | 252 | 256 | 182.69 | 252.7332 | 245 | 253.1783 | 250 | 239.5094 | 257 | 249 | 264 | 226 | 66
| 1 | 6 | 249 | 250 | 225.96 | 246.9135 | 255.3413 | 168 | 310 | 249 | 239 | 241 | 260 | 246 | 259 | 246 | 253 | 257 | 250 | 245 | 271 | 182.43 | 251.4268 | 250 | 253.3234 | 250 | 229.4708 | 255 | 254 | 266 | 255
| Mean |  | 257.8333 | 249.6667 | 225.0767 | 244.3497 | 251.9527 | 166.8333 | 310 | 243.1667 | 242.5 | 240.8333 | 258.8333 | 251 | 260.3333 | 251.5 | 248.6667 | 252.3333 | 256.6667 | 247.3333 | 262.1667 | 198.0667 | 254.8063 | 248.3333 | 253.4247 | 249.1667 | 233.5541 | 253.8333 | 246.8333 | 265.3333 | 238
| Median |  | 258 | 250 | 224.77 | 244.5276 | 249.4896 | 166.5 | 311.5 | 243 | 241.5 | 241 | 259.5 | 248.5 | 259.5 | 248 | 249 | 255 | 256.5 | 247.5 | 261 | 202.75 | 254.1178 | 250 | 253.4794 | 250 | 232.3085 | 254.5 | 245.5 | 265 | 239.5
| Std Dev. |  | 5.5287 | 2.7325 | 1.388 | 3.5058 | 4.9951 | 20.6728 | 6.603 | 5.7067 | 7.1204 | 0.75277 | 3.0605 | 5.933 | 7.6855 | 11.1848 | 3.266 | 7.6333 | 9.0701 | 3.266 | 5.6362 | 12.4091 | 3.2234 | 2.582 | 0.37423 | 3.6009 | 3.8458 | 3.4303 | 4.0702 | 1.5055 | 14.0143
| Rel.Std.Dev. |  | 2.144% | 1.094% | 0.6167% | 1.435% | 1.983% | 12.39% | 2.13% | 2.347% | 2.936% | 0.3126% | 1.182% | 2.364% | 2.952% | 4.447% | 1.313% | 3.025% | 3.534% | 1.32% | 2.15% | 6.265% | 1.265% | 1.04% | 0.1477% | 1.445% | 1.647% | 1.351% | 1.649% | 0.5674% | 5.888%
| PDM3 |  | 2.943% | -0.3181% | -10.14% | -2.441% | 0.5946% | -33.39% | 23.77% | -2.913% | -3.179% | -3.845% | 3.342% | 0.2142% | 3.941% | 0.4139% | -0.7174% | 0.7466% | 2.477% | -1.25% | 4.673% | -20.92% | 1.734% | -0.8505% | 1.182% | -0.5177% | -6.751% | 1.345% | -1.449% | 5.937% | -4.976%
| Z-Score (Absolute) |  | 0.85 | 0 | 2.55 | 0.55 | 0.24 | 8.59 | 6.26 | 0.67 | 0.74 | 0.92 | 0.95 | 0.14 | 1.11 | 0.19 | 0.1 | 0.28 | 0.73 | 0.24 | 1.3 | 5.35 | 0.53 | 0.14 | 0.39 | 0.05 | 1.67 | 0.43 | 0.29 | 1.63 | 1.21
| Analytical results for W in OREAS 241b (Certified Value 38.3 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 40.9 | 37.9 | 38.93 | 29.69 | 38.9498 | 2.4 | 45.2 | 34.8 | 36.8 | 38 | 37 | 40.2 | 40.4 | 35.3 | 39.2 | 32 | 33 | 34.1 | 37 | 46 | 37.8355 | 40 | 39.1084 | 41.3 | 37.8 | 39.6 | 41.3 |  |  | 1
| 1 | 2 | 41.5 | 38.8 | 40.66 | 29.19 | 40.082 | 3.8 | 44.2 | 36.1 | 38.1 | 38 | 37.6 | 40.5 | 38.8 | 35.5 | 37.6 | 32 | 36.3 | 33.6 | 40 | 45 | 37.8034 | 37.4 | 39.4297 | 40.7 | 36.7 | 40.5 | 40.5 |  |  | 30
| 1 | 3 | 41.7 | 37.3 | 39.13 | 29.42 | 39.5541 | 1.2 | 40.7 | 36.2 | 37.7 | 37 | 37.7 | 40 | 38.3 | 35 | 39 | 32 | 37.3 | 33.9 | 38 | 47 | 38.7122 | 40.2 | 39.6554 | 41 | 36.8 | 39.8 | 39.3 |  |  | 16
| 1 | 4 | 42 | 37.9 | 42.2 | 28.73 | 38.6177 | 1.6 | 40.4 | 35.2 | 37.1 | 38 | 37.9 | 39.7 | 40.8 | 34.8 | 37.1 | 34 | 33.5 | 33.5 | 38 | 45 | 38.2932 | 38.7 | 39.5682 | 40.7 | 36.8 | 39 | 39.7 |  |  | 38.2764
| 1 | 5 | 44.6 | 38 | 39.49 | 29.35 | 39.2765 | 3.8 | 40 | 35.9 | 37.4 | 38 | 34.1 | 40.6 | 42.3 | 36.2 | 37.6 | 33 | 36.2 | 35 | 37 | 44 | 38.3646 | 41.3 | 39.0164 | 39.4 | 38.8 | 40.4 | 42.1 |  |  | 67
| 1 | 6 | 42 | 37.5 | 41.07 | 29.82 | 40.5585 | 1.2 | 38.7 | 35.7 | 37.9 | 37 | 35 | 39.6 | 39.8 | 35.8 | 36.5 | 34 | 35.7 | 34.3 | 40 | 46 | 38.4283 | 39 | 39.6128 | 40.5 | 38.6 | 39.9 | 43.7
| Mean |  | 42.1167 | 37.9 | 40.2467 | 29.3667 | 39.5064 | 2.3333 | 41.5333 | 35.65 | 37.5 | 37.6667 | 36.55 | 40.1 | 40.0667 | 35.4333 | 37.8333 | 32.8333 | 35.3333 | 34.0667 | 38.3333 | 45.5 | 38.2395 | 39.4333 | 39.3985 | 40.6 | 37.5833 | 39.8667 | 41.1
| Median |  | 41.85 | 37.9 | 40.075 | 29.385 | 39.4153 | 2 | 40.55 | 35.8 | 37.55 | 38 | 37.3 | 40.1 | 40.1 | 35.4 | 37.6 | 32.5 | 35.95 | 34 | 38 | 45.5 | 38.3289 | 39.5 | 39.4989 | 40.7 | 37.3 | 39.85 | 40.9
| Std Dev. |  | 1.2828 | 0.51769 | 1.282 | 0.38723 | 0.72012 | 1.2176 | 2.5657 | 0.54681 | 0.49396 | 0.5164 | 1.6034 | 0.40988 | 1.4445 | 0.5164 | 1.0633 | 0.98319 | 1.7025 | 0.5465 | 1.3663 | 1.0488 | 0.35524 | 1.3604 | 0.27271 | 0.65115 | 0.95586 | 0.55015 | 1.6346
| Rel.Std.Dev. |  | 3.046% | 1.366% | 3.185% | 1.319% | 1.823% | 52.18% | 6.177% | 1.534% | 1.317% | 1.371% | 4.387% | 1.022% | 3.605% | 1.457% | 2.811% | 2.994% | 4.819% | 1.604% | 3.564% | 2.305% | 0.929% | 3.45% | 0.6922% | 1.604% | 2.543% | 1.38% | 3.977%
| PDM3 |  | 10.03% | -0.9835% | 5.147% | -23.28% | 3.213% | -93.9% | 8.509% | -6.862% | -2.028% | -1.593% | -4.51% | 4.764% | 4.677% | -7.428% | -1.158% | -14.22% | -7.689% | -11% | 0.1486% | 18.87% | -0.09639% | 3.022% | 2.931% | 6.07% | -1.811% | 4.155% | 7.377%
| Z-Score (Absolute) |  | 1.41 | 0.12 | 0.73 | 3.22 | 0.46 | 13.01 | 1.19 | 0.94 | 0.27 | 0.21 | 0.61 | 0.67 | 0.66 | 1.02 | 0.15 | 1.96 | 1.05 | 1.51 | 0.03 | 2.63 | 0 | 0.43 | 0.42 | 0.86 | 0.24 | 0.59 | 1.04
| Analytical results for Y in OREAS 241b (Certified Value 19.5 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
2
| 1 | 1 | 18.3 | 19.16 | 17.15 | 17.1912 | 19.5116 | 18.7 | 23.1 | 19.9 | 18.31 | 19 | 20.5 | 20.2 | 21.2 | 20.6 | 20.7 | 18 | 18.1 | 18.5 | 19.7 | 23 | 20.3787 | 19.7 | 18.2977 | 20.4 | 18.8 | 20.4 | 19.8 |  |  | 1
| 1 | 2 | 18.7 | 19.46 | 17.34 | 17.943 | 20.1623 | 18 | 22.7 | 20 | 19.16 | 19 | 20.8 | 20.5 | 20.2 | 20.3 | 21.2 | 17 | 19.3 | 18 | 20 | 23.1 | 20.4473 | 18.7 | 18.3848 | 20.5 | 19.2 | 20.5 | 19.1 |  |  | 31
| 1 | 3 | 18.9 | 19.64 | 17.37 | 17.6321 | 19.2987 | 18.3 | 21.1 | 20.1 | 18.95 | 18 | 20.7 | 20 | 20.6 | 20 | 21 | 18 | 19.8 | 18.9 | 19.7 | 23.8 | 20.6422 | 20.4 | 18.3382 | 20.5 | 19.4 | 20.2 | 18.7 |  |  | 16
| 1 | 4 | 18.6 | 19.38 | 17.17 | 17.4255 | 19.4492 | 18.8 | 21.1 | 19.8 | 19.21 | 19 | 20.8 | 20.3 | 21.3 | 19.6 | 20 | 19 | 18.3 | 18.1 | 18.8 | 23.5 | 20.0191 | 20.1 | 18.2758 | 20 | 19 | 20.2 | 19.3 |  |  | 19.5048
| 1 | 5 | 19.4 | 19.67 | 17.48 | 17.9738 | 19.4926 | 18.1 | 20.8 | 20.4 | 19.33 | 18 | 20.7 | 20.3 | 22 | 20.5 | 21.3 | 19 | 19.9 | 18.9 | 20.1 | 22.1 | 20.3957 | 20.4 | 17.8914 | 20.7 | 19.6 | 20.8 | 20.1 |  |  | 68
| 1 | 6 | 19.1 | 19.18 | 17.06 | 17.6653 | 19.9484 | 19.1 | 19.8 | 20.8 | 19.39 | 18 | 21.4 | 20.1 | 21.7 | 19.5 | 20.4 | 19 | 19.5 | 18.2 | 20.1 | 23.8 | 20.2434 | 18.8 | 18.2145 | 20.8 | 19.3 | 20.3 | 21.2
| Mean |  | 18.8333 | 19.415 | 17.2617 | 17.6385 | 19.6438 | 18.5 | 21.4333 | 20.1667 | 19.0583 | 18.5 | 20.8167 | 20.2333 | 21.1667 | 20.0833 | 20.7667 | 18.3333 | 19.15 | 18.4333 | 19.7333 | 23.2167 | 20.3544 | 19.6833 | 18.2337 | 20.4833 | 19.2167 | 20.4 | 19.7
| Median |  | 18.8 | 19.42 | 17.255 | 17.6487 | 19.5021 | 18.5 | 21.1 | 20.05 | 19.185 | 18.5 | 20.75 | 20.25 | 21.25 | 20.15 | 20.85 | 18.5 | 19.4 | 18.35 | 19.85 | 23.3 | 20.3872 | 19.9 | 18.2867 | 20.5 | 19.25 | 20.35 | 19.55
| Std Dev. |  | 0.38816 | 0.2187 | 0.15943 | 0.30044 | 0.33435 | 0.43359 | 1.2388 | 0.37238 | 0.39721 | 0.54772 | 0.30605 | 0.17512 | 0.67132 | 0.46224 | 0.50067 | 0.8165 | 0.76877 | 0.39833 | 0.49261 | 0.64317 | 0.20903 | 0.7679 | 0.17729 | 0.27869 | 0.28577 | 0.22804 | 0.88769
| Rel.Std.Dev. |  | 2.061% | 1.126% | 0.9236% | 1.703% | 1.702% | 2.344% | 5.78% | 1.847% | 2.084% | 2.961% | 1.47% | 0.8655% | 3.172% | 2.302% | 2.411% | 4.454% | 4.014% | 2.161% | 2.496% | 2.77% | 1.027% | 3.901% | 0.9723% | 1.361% | 1.487% | 1.118% | 4.506%
| PDM3 |  | -3.443% | -0.4605% | -11.5% | -9.569% | 0.7126% | -5.152% | 9.887% | 3.393% | -2.289% | -5.152% | 6.726% | 3.735% | 8.52% | 2.966% | 6.469% | -6.006% | -1.819% | -5.493% | 1.172% | 19.03% | 4.356% | 0.9153% | -6.517% | 5.017% | -1.477% | 4.59% | 1.001%
| Z-Score (Absolute) |  | 0.72 | 0.23 | 2.04 | 1.72 | 0.03 | 1 | 1.48 | 0.41 | 0.53 | 1 | 0.96 | 0.46 | 1.25 | 0.34 | 0.91 | 1.14 | 0.45 | 1.05 | 0.04 | 2.98 | 0.57 | 0 | 1.22 | 0.67 | 0.39 | 0.6 | 0.01
| Analytical results for Yb in OREAS 241b (Certified Value 2.1 ppm) | [30 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.24 | 1.27 | 1.28 | 1.32 | [19 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | [19 empty] | 2
3
| 1 | 1 | 2.3 | 2.04 | 1.85 | 1.9449 | 2.1 | 2.35 | 2.2 | 2 | 2.2259 | 2 | [19 empty] | 1
| 1 | 2 | 2.4 | 2.1 | 1.78 | 1.9788 | 2.1 | 2.33 | 2.2 | 2.1 | 2.2704 | 2 | [19 empty] | 32
| 1 | 3 | 2.4 | 2.16 | 1.76 | 1.9673 | 2 | 2.15 | 2.2 | 2 | 2.2862 | 2 | [19 empty] | 16
| 1 | 4 | 2.3 | 2.16 | 1.81 | 1.9392 | 2.1 | 2.16 | 2.2 | 2.1 | 2.2799 | 2 | [19 empty] | 2.1017
| 1 | 5 | 2.4 | 2.03 | 1.81 | 1.9539 | 2.1 | 2.12 | 2.1 | 2 | 2.2547 | 2 | [19 empty] | 69
| 1 | 6 | 2.4 | 2.08 | 1.75 | 1.9308 | 2.2 | 2.04 | 2.2 | 2.1 | 2.2895 | 2.1
| Mean |  | 2.3667 | 2.095 | 1.7933 | 1.9525 | 2.1 | 2.1917 | 2.1833 | 2.05 | 2.2678 | 2.0167
| Median |  | 2.4 | 2.09 | 1.795 | 1.9494 | 2.1 | 2.155 | 2.2 | 2.05 | 2.2752 | 2
| Std Dev. |  | 0.05164 | 0.05648 | 0.037238 | 0.017994 | 0.063246 | 0.12254 | 0.040825 | 0.054772 | 0.024041 | 0.040825
| Rel.Std.Dev. |  | 2.182% | 2.696% | 2.076% | 0.9216% | 3.012% | 5.591% | 1.87% | 2.672% | 1.06% | 2.024%
| PDM3 |  | 12.61% | -0.3184% | -14.67% | -7.099% | -0.08046% | 4.281% | 3.885% | -2.459% | 7.902% | -4.046%
| Z-Score (Absolute) |  | 2.02 | 0.02 | 2.28 | 1.09 | 0.02 | 0.71 | 0.64 | 0.36 | 1.28 | 0.61
| Analytical results for Zn in OREAS 241b (Certified Value 163 ppm) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 163 | 169 | 156.76 | 159.49 | 164.69 | 157 | 211 | 162 | 161 | 153 | 164 | 166 | 190 | 171 | 173 | 158 | 175 | 157 | 161 | 72.14 | 157.3302 | 160 | 161.3996 | 170 | 193.9122 | 150 | 164 | 153 | 166 | 1
| 1 | 2 | 164 | 171 | 158.01 | 164.51 | 169.5958 | 157 | 204 | 164 | 164 | 150 | 162 | 167 | 180 | 169 | 171 | 152 | 172 | 155 | 163 | 72.8 | 153.3172 | 160 | 159.1232 | 168 | 196.9372 | 149 | 158 | 151 | 159 | 33
| 1 | 3 | 165 | 172 | 155.28 | 163.23 | 163.3588 | 158 | 185 | 167 | 163 | 151 | 164 | 171 | 179 | 168 | 173 | 156 | 174 | 156 | 159 | 73.8 | 154.653 | 160 | 163.7912 | 173 | 196.266 | 154 | 161 | 154 | 153 | 16
| 1 | 4 | 169 | 167 | 155.28 | 161.34 | 164.9868 | 161 | 198 | 164 | 161 | 153 | 166 | 175 | 182 | 166 | 172 | 164 | 171 | 157 | 159 | 67.1 | 153.6927 | 160 | 160.86 | 169 | 204.0162 | 154 | 163 | 153 | 160 | 162.9386
| 1 | 5 | 164 | 169 | 156.72 | 170.26 | 164.4935 | 161 | 181 | 167 | 165 | 148 | 163 | 167 | 187 | 184 | 172 | 162 | 172 | 160 | 157 | 72.13 | 151.8885 | 160 | 160.4752 | 174 | 193.6752 | 151 | 165 | 152 | 168 | 70
| 1 | 6 | 165 | 168 | 155.8 | 162.45 | 168.3705 | 158 | 172 | 168 | 165 | 150 | 159 | 163 | 182 | 169 | 174 | 163 | 175 | 158 | 164 | 90.8 | 153.3435 | 160 | 159.8814 | 173 | 193.8582 | 154 | 165 | 153 | 173
| Mean |  | 165 | 169.3333 | 156.3083 | 163.5467 | 165.9159 | 158.6667 | 191.8333 | 165.3333 | 163.1667 | 150.8333 | 163 | 168.1667 | 183.3333 | 171.1667 | 172.5 | 159.1667 | 173.1667 | 157.1667 | 160.5 | 74.795 | 154.0375 | 160 | 160.9218 | 171.1667 | 196.4442 | 152 | 162.6667 | 152.6667 | 163.1667
| Median |  | 164.5 | 169 | 156.26 | 162.84 | 164.8384 | 158 | 191.5 | 165.5 | 163.5 | 150.5 | 163.5 | 167 | 182 | 169 | 172.5 | 160 | 173 | 157 | 160 | 72.47 | 153.5181 | 160 | 160.6676 | 171.5 | 195.0891 | 152.5 | 163.5 | 153 | 163
| Std Dev. |  | 2.0976 | 1.8619 | 1.0623 | 3.7047 | 2.4696 | 1.8619 | 14.9053 | 2.3381 | 1.8348 | 1.9408 | 2.3664 | 4.2151 | 4.274 | 6.4936 | 1.0488 | 4.6655 | 1.7224 | 1.7224 | 2.6646 | 8.1792 | 1.8416 | 0 | 1.6114 | 2.4833 | 3.9592 | 2.2804 | 2.7325 | 1.0328 | 7.1949
| Rel.Std.Dev. |  | 1.271% | 1.1% | 0.6796% | 2.265% | 1.488% | 1.173% | 7.77% | 1.414% | 1.125% | 1.287% | 1.452% | 2.506% | 2.331% | 3.794% | 0.608% | 2.931% | 0.9946% | 1.096% | 1.66% | 10.94% | 1.196% | 0% | 1.001% | 1.451% | 2.015% | 1.5% | 1.68% | 0.6765% | 4.41%
| PDM3 |  | 1.265% | 3.925% | -4.069% | 0.3732% | 1.827% | -2.622% | 17.73% | 1.47% | 0.14% | -7.429% | 0.03767% | 3.209% | 12.52% | 5.05% | 5.868% | -2.315% | 6.277% | -3.542% | -1.497% | -54.1% | -5.463% | -1.804% | -1.238% | 5.05% | 20.56% | -6.713% | -0.1669% | -6.304% | 0.14%
| Z-Score (Absolute) |  | 0.21 | 0.69 | 0.77 | 0.04 | 0.31 | 0.51 | 3.22 | 0.24 | 0 | 1.39 | 0.02 | 0.56 | 2.27 | 0.9 | 1.05 | 0.45 | 1.12 | 0.67 | 0.3 | 9.93 | 1.03 | 0.36 | 0.25 | 0.9 | 3.74 | 1.25 | 0.06 | 1.18 | 0
| Analytical results for Zr in OREAS 241b (Certified Value 57 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
1
| 1 | 1 | 61.6 | 55.8 | 71.17 | 61.3234 | 59.5243 | 40 | 66 | 54.2 | 49.9 | 61 | 62.9 | 55.3 | 57.7 | 67.5 | 49 | 48 | 56.8 | 51.3 | 53 | 58.0428 | 55 | 78.2691 | 59.3 | 34.9033 | 54.2 | 87.4 | 55.6 | 51.7 |  | 1
| 1 | 2 | 64.7 | 54.4 | 71.86 | 62.5941 | 60.5417 | 48 | 67 | 55.4 | 52.3 | 61 | 62.1 | 55.7 | 58.1 | 57.4 | 54.2 | 43 | 57.9 | 49.8 | 53.9 | 55.89 | 55 | 79.1166 | 51.1 | 35.221 | 53.7 | 85.3 | 56.8 | 49.6 |  | 34
| 1 | 3 | 65.7 | 55.5 | 75.57 | 61.886 | 57.5587 | 28 | 68 | 56 | 51.6 | 60 | 60.2 | 56.5 | 62 | 62.9 | 51 | 45 | 57.5 | 50.2 | 51.9 | 55.5588 | 60 | 77.4795 | 65.2 | 35.2512 | 55.5 | 85.9 | 54.7 | 48.9 |  | 16
| 1 | 4 | 59.5 | 53 | 77.22 | 62.9627 | 59.1192 | 30 | 63 | 55.4 | 54 | 61 | 63.1 | 60.2 | 62 | 67.7 | 52.2 | 49 | 56.9 | 53.4 | 51.5 | 55.522 | 55 | 76.3487 | 61 | 34.7688 | 54 | 87.1 | 53.5 | 51.4 |  | 57.3284
| 1 | 5 | 64.7 | 57.9 | 70.72 | 65.7854 | 58.8411 | 41 | 64 | 56.6 | 53.4 | 62 | 63.4 | 56.3 | 59.6 | 59.8 | 51 | 49 | 57.8 | 50.4 | 50.6 | 54.8044 | 55 | 78.2268 | 63 | 33.957 | 56.4 | 86.5 | 57.1 | 52.5 |  | 71
| 1 | 6 | 62.2 | 55.4 | 68.97 | 63.9424 | 58.0933 | 39 | 65 | 56.5 | 52.3 | 59 | 63.7 | 54.9 | 64.8 | 55 | 65.2 | 48 | 57.4 | 49.5 | 54.6 | 55.4208 | 55 | 76.7789 | 52.3 | 34.7477 | 54.6 | 88 | 55.1 | 54.9
| Mean |  | 63.0667 | 55.3333 | 72.585 | 63.0823 | 58.9464 | 37.6667 | 65.5 | 55.6833 | 52.25 | 60.6667 | 62.5667 | 56.4833 | 60.7 | 61.7167 | 53.7667 | 47 | 57.3833 | 50.7667 | 52.5833 | 55.8731 | 55.8333 | 77.7033 | 58.65 | 34.8082 | 54.7333 | 86.7 | 55.4667 | 51.5
| Median |  | 63.45 | 55.45 | 71.515 | 62.7784 | 58.9801 | 39.5 | 65.5 | 55.7 | 52.3 | 61 | 63 | 56 | 60.8 | 61.35 | 51.6 | 48 | 57.45 | 50.3 | 52.45 | 55.5404 | 55 | 77.8532 | 60.15 | 34.8361 | 54.4 | 86.8 | 55.35 | 51.55
| Std Dev. |  | 2.3619 | 1.6219 | 3.1454 | 1.6015 | 1.0558 | 7.4476 | 1.8708 | 0.89088 | 1.4377 | 1.0328 | 1.2801 | 1.9167 | 2.7262 | 5.2549 | 5.8548 | 2.4495 | 0.4535 | 1.4292 | 1.5198 | 1.1204 | 2.0412 | 1.0326 | 5.7462 | 0.4702 | 1.027 | 0.998 | 1.3456 | 2.1438
| Rel.Std.Dev. |  | 3.745% | 2.931% | 4.333% | 2.539% | 1.791% | 19.77% | 2.856% | 1.6% | 2.752% | 1.702% | 2.046% | 3.393% | 4.491% | 8.515% | 10.89% | 5.212% | 0.7903% | 2.815% | 2.89% | 2.005% | 3.656% | 1.329% | 9.797% | 1.351% | 1.876% | 1.151% | 2.426% | 4.163%
| PDM3 |  | 10.01% | -3.48% | 26.61% | 10.04% | 2.822% | -34.3% | 14.25% | -2.87% | -8.858% | 5.823% | 9.137% | -1.474% | 5.881% | 7.655% | -6.213% | -18.02% | 0.09585% | -11.45% | -8.277% | -2.538% | -2.608% | 35.54% | 2.305% | -39.28% | -4.527% | 51.23% | -3.247% | -10.17%
| Z-Score (Absolute) |  | 1.03 | 0.13 | 2.46 | 1.03 | 0.41 | 2.77 | 1.4 | 0.07 | 0.59 | 0.67 | 0.96 | 0.05 | 0.68 | 0.83 | 0.36 | 1.37 | 0.18 | 0.81 | 0.54 | 0.05 | 0.05 | 3.22 | 0.37 | 3.2 | 0.22 | 4.57 | 0.11 | 0.7
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C256: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C274: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C292: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C310: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C328: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C346: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C365: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C383: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C402: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C421: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C439: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C513: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C551: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C569: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C606: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C624: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C661: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C679: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C697: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C715: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C733: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C751: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C769: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C788: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C897: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C916: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1009: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1028: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1100: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1118: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1136: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 241b (Certified Value 1.73 ppm) | [32 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 2.39 | 1.72 | 1.86 | 1.7396 | 1.69 | 1.88 | 1.76 | 1.7 | 1.8 | 1.8 | 1.78 | 1.73 | 1.68 | < 10 | 1.72 | 1.6 | 1.1822 | 1.65 | 1.586 | 1.77 | 1.75 | 1.4 | 1.7728 | 1.8 | 1.7 | 1.64 | 1.81 | 1.75 | 1.81 |  |  | 1
| 1 | 2 | 2.45 | 1.73 | 1.86 | 1.7623 | 1.73 | 1.88 | 1.68 | 1.81 | 1.9 | 1.7 | 1.84 | 1.78 | 1.67 | < 10 | 1.74 | 1.6 | < 1 | 1.57 | 1.658 | 1.72 | 1.32 | 1.4 | 1.8033 | 1.8 | 1.889 | 1.73 | 1.71 | 1.75 | 1.78 |  |  | 18
| 1 | 3 | 2.42 | 1.74 | 3.26 | 1.7184 | 1.77 | 1.86 | 1.74 | 1.63 | 1.8 | 1.8 | 1.88 | 1.73 | 1.72 | < 10 | 1.72 | 1.6 | 1.5351 | 1.65 | 1.636 | 1.75 | 1.48 | 1.4 | 1.7613 | 1.82 | 1.845 | 1.76 | 1.75 | 1.74 | 1.87 |  |  | 16
| 1 | 4 | 2.44 | 1.74 | 2.07 | 1.7508 | 1.66 | 1.83 | 1.69 | 1.68 | 1.8 | 1.7 | 1.8 | 1.69 | 1.71 | < 10 | 1.69 | 1.6 | 1.9249 | 1.65 | 1.57 | 1.82 | 1.87 | 1.4 | 1.7348 | 1.74 | 1.847 | 1.7 | 1.78 | 1.74 | 1.78 |  |  | 1.7333
| 1 | 5 | 2.39 | 1.72 | 1.72 | 1.7531 | 1.71 | 1.83 | 1.75 | 1.69 | 1.8 | 1.8 | 1.91 | 1.71 | 1.68 | < 10 | 1.71 | 1.6 | 1.6469 | 1.58 | 1.626 | 1.78 | 1.43 | 1.4 | 1.7392 | 1.83 | 1.777 | 1.7 | 1.8 | 1.73 | 1.86 |  |  | 73
| 1 | 6 | 2.48 | 1.72 | 1.83 | 1.7134 | 1.62 | 1.89 | 1.75 | 1.67 | 1.9 | 1.8 | 1.76 | 1.7 | 1.72 | < 10 | 1.71 | 1.6 | 1.657 | 1.6 | 1.557 | 1.69 | 1.58 | 1.5 | 1.7456 | 1.73 | 1.769 | 1.66 | 1.71 | 1.73 | 1.81
| Mean |  | 2.4283 | 1.7283 | 2.1 | 1.7396 | 1.6967 | 1.8617 | 1.7283 | 1.6967 | 1.8333 | 1.7667 | 1.8283 | 1.7233 | 1.6967 |  | 1.715 | 1.6 | 1.5892 | 1.6167 | 1.6055 | 1.755 | 1.5717 | 1.4167 | 1.7595 | 1.7867 | 1.8045 | 1.6983 | 1.76 | 1.74 | 1.8183
| Median |  | 2.43 | 1.725 | 1.86 | 1.7452 | 1.7 | 1.87 | 1.745 | 1.685 | 1.8 | 1.8 | 1.82 | 1.72 | 1.695 |  | 1.715 | 1.6 | 1.6469 | 1.625 | 1.606 | 1.76 | 1.53 | 1.4 | 1.7534 | 1.8 | 1.811 | 1.7 | 1.765 | 1.74 | 1.81
| Std Dev. |  | 0.035449 | 0.0098319 | 0.57948 | 0.019791 | 0.052789 | 0.026394 | 0.034303 | 0.060553 | 0.05164 | 0.05164 | 0.058793 | 0.032042 | 0.022509 |  | 0.016432 | 2.4324e-16 | 0.26889 | 0.037771 | 0.040248 | 0.045935 | 0.20624 | 0.040825 | 0.025745 | 0.041793 | 0.068591 | 0.044008 | 0.043818 | 0.0089443 | 0.038687
| Rel.Std.Dev. |  | 1.46% | 0.5689% | 27.59% | 1.138% | 3.111% | 1.418% | 1.985% | 3.569% | 2.817% | 2.923% | 3.216% | 1.859% | 1.327% |  | 0.9581% | 1.52e-14% | 16.92% | 2.336% | 2.507% | 2.617% | 13.12% | 2.882% | 1.463% | 2.339% | 3.801% | 2.591% | 2.49% | 0.514% | 2.128%
| PDM3 |  | 40.1% | -0.289% | 21.15% | 0.3614% | -2.116% | 7.403% | -0.289% | -2.116% | 5.769% | 1.923% | 5.48% | -0.5774% | -2.116% |  | -1.058% | -7.693% | -8.316% | -6.731% | -7.375% | 1.25% | -9.327% | -18.27% | 1.509% | 3.076% | 4.105% | -2.02% | 1.538% | 0.3841% | 4.903%
| Z-Score (Absolute) |  | 9.87 | 0.16 | 5.16 | 0 | 0.62 | 1.75 | 0.16 | 0.62 | 1.34 | 0.39 | 1.27 | 0.23 | 0.62 | 46.72 | 0.35 | 2 | 4.76 | 1.76 | 1.92 | 0.22 | 2.41 | 4.63 | 0.28 | 0.67 | 0.93 | 0.59 | 0.29 | 0.01 | 1.13
| Analytical results for Al in OREAS 241b (Certified Value 3.04 wt.%) | [32 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 3
| 1 | 1 | 3.1 | 2.862 | 2.6141 | 3.409 | 2.9498 | 2.85 | 3.281 | 3.47 | 2.8865 | 2.83 | 3.15 | 3.14 | 3.03 | 3.03 | 3.07 | 3.01 | 3.2068 | 3.482 | 2.79 | 3.35 | 3.3875 | 2.8304 | 2.6265 | 2.99 | 2.75 | 3.3 | 3.06 | 2.8 |  |  |  | 1
| 1 | 2 | 3.07 | 2.848 | 2.6731 | 3.4269 | 3.0735 | 3 | 3.345 | 3.4 | 3.019 | 2.83 | 3.16 | 3.04 | 3.03 | 3.05 | 3.1 | 3.03 | 3.3398 | 3.394 | 3.01 | 3.31 | 3.6014 | 2.7945 | 2.6735 | 2.99 | 2.88 | 3.3 | 2.97 | 2.75 |  |  |  | #N/A
| 1 | 3 | 3.06 | 2.812 | 2.6911 | 3.4441 | 3.0497 | 2.95 | 3.302 | 3.42 | 2.8764 | 2.84 | 3.15 | 3.07 | 2.99 | 3.05 | 3.07 | 3.01 | 3.2652 | 3.379 | 2.96 | 3.34 | 3.5094 | 2.7638 | 2.7535 | 3 | 2.9 | 3.22 | 2.97 | 2.83 |  |  |  | 16
| 1 | 4 | 3.07 | 2.806 | 2.5952 | 3.4225 | 2.9953 | 2.91 | 3.308 | 3.39 | 2.8778 | 2.87 | 3.16 | 3.05 | 3.05 | 3.04 | 3.06 | 3.04 | 3.3096 | 3.376 | 2.72 | 3.33 | 3.1096 | 2.698 | 2.6325 | 2.95 | 2.9 | 3.3 | 2.99 | 2.71 |  |  |  | 3.044
| 1 | 5 | 3.13 | 2.835 | 2.6803 | 3.4547 | 3.0895 | 2.85 | 3.371 | 3.45 | 2.75 | 2.89 | 3.2 | 3.04 | 3.04 | 3.08 | 2.99 | 2.99 | 3.0668 | 3.472 | 3.03 | 3.36 | 2.9372 | 2.6083 | 2.6725 | 3 | 2.84 | 3.26 | 2.99 | 2.85 |  |  |  | 74
| 1 | 6 | 3.05 | 2.815 | 2.6297 | 3.4387 | 2.9388 | 2.67 | 3.345 | 3.52 | 2.7463 | 2.87 | 3.18 | 3.06 | 3.04 | 3.06 | 3.05 | 2.98 | 3.1668 | 3.388 | 2.91 | 3.3 | 3.2448 | 2.8745 | 2.7765 | 2.99 | 2.69 | 3.25 | 2.95 | 2.82
| Mean |  | 3.08 | 2.8297 | 2.6473 | 3.4327 | 3.0161 | 2.8717 | 3.3253 | 3.4417 | 2.8593 | 2.855 | 3.1667 | 3.0667 | 3.03 | 3.0517 | 3.0567 | 3.01 | 3.2258 | 3.4152 | 2.9033 | 3.3317 | 3.2983 | 2.7616 | 2.6892 | 2.9867 | 2.8267 | 3.2717 | 2.9883 | 2.7933
| Median |  | 3.07 | 2.825 | 2.6514 | 3.4328 | 3.0225 | 2.88 | 3.3265 | 3.435 | 2.8771 | 2.855 | 3.16 | 3.055 | 3.035 | 3.05 | 3.065 | 3.01 | 3.236 | 3.391 | 2.935 | 3.335 | 3.3161 | 2.7792 | 2.673 | 2.99 | 2.86 | 3.28 | 2.98 | 2.81
| Std Dev. |  | 0.029665 | 0.022349 | 0.03948 | 0.016428 | 0.064221 | 0.11462 | 0.033673 | 0.048751 | 0.10158 | 0.0251 | 0.019664 | 0.037771 | 0.020976 | 0.017224 | 0.036697 | 0.022804 | 0.10068 | 0.048425 | 0.1242 | 0.023166 | 0.2503 | 0.096054 | 0.062334 | 0.018619 | 0.08756 | 0.033714 | 0.038166 | 0.053166
| Rel.Std.Dev. |  | 0.9631% | 0.7898% | 1.491% | 0.4786% | 2.129% | 3.991% | 1.013% | 1.416% | 3.553% | 0.8792% | 0.621% | 1.232% | 0.6923% | 0.5644% | 1.201% | 0.7576% | 3.121% | 1.418% | 4.278% | 0.6953% | 7.589% | 3.478% | 2.318% | 0.6234% | 3.098% | 1.03% | 1.277% | 1.903%
| PDM3 |  | 1.182% | -7.041% | -13.03% | 12.77% | -0.917% | -5.662% | 9.242% | 13.06% | -6.067% | -6.209% | 4.03% | 0.7444% | -0.4601% | 0.2517% | 0.4159% | -1.117% | 5.974% | 12.19% | -4.621% | 9.45% | 8.355% | -9.278% | -11.66% | -1.884% | -7.14% | 7.479% | -1.829% | -8.235%
| Z-Score (Absolute) |  | 0.21 | 0.72 | 1.4 | 1.53 | 0.03 | 0.56 | 1.13 | 1.56 | 0.61 | 0.63 | 0.54 | 0.16 | 0.03 | 0.11 | 0.13 | 0.05 | 0.76 | 1.46 | 0.45 | 1.15 | 1.03 | 0.98 | 1.25 | 0.14 | 0.73 | 0.93 | 0.13 | 0.86
| Analytical results for As in OREAS 241b (Certified Value 86 ppm) | [32 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 1
| 1 | 1 | 90 | 84.71 | 73.66 | 80.3042 | 83.9986 | 82.9 | 86.8 | 89 | 85 | 78 | 81.3 | 90.6 | 89.3 | 83.7 | < 100 | 89 | 79 | 74.5678 | 86 | 85.1 | 89 | 61.77 | 89.9852 | 80 | 90.3409 | 88.4 | 77.162 | 87.7 | 89 | 87.8 | 86.1 | 1
| 1 | 2 | 89 | 85.38 | 75.99 | 80.7272 | 84.8822 | 87 | 88.7 | 87 | 89 | 78 | 80.6 | 90.5 | 86.8 | 85.3 | < 100 | 89.4 | 79 | 72.6052 | 87 | 87.5 | 84 | 61.16 | 91.034 | 85 | 91.8012 | 88.8 | 90.475 | 87.6 | 84 | 85.8 | 85.4 | 19
| 1 | 3 | 90 | 84 | 76.77 | 79.5992 | 82.0974 | 86.9 | 89 | 89 | 84 | 78 | 83.6 | 91.3 | 87.8 | 85.4 | < 100 | 88.8 | 78 | 70.5648 | 90 | 88.5 | 84 | 61.94 | 89.102 | 85 | 87.0402 | 89.1 | 81.801 | 87.1 | 88 | 83 | 88 | 16
| 1 | 4 | 90 | 85.61 | 75.34 | 80.7554 | 85.0714 | 82.9 | 90.5 | 87 | 84 | 79 | 83.3 | 89.1 | 89.4 | 85.8 | < 100 | 88.2 | 81 | 73.2766 | 85 | 82.2 | 84 | 62.33 | 88.1084 | 85 | 91.2618 | 86.3 | 84.494 | 87 | 89 | 84.6 | 85.1 | 85.5163
| 1 | 5 | 91 | 84.6 | 76.12 | 81.0374 | 85.7075 | 83.9 | 86.2 | 89 | 83 | 79 | 82.3 | 90.8 | 89.5 | 85.4 | 104.113 | 87 | 79 | 74.3513 | 89 | 85 | 85 | 62.28 | 80.9692 | 85 | 91.4905 | 90.4 | 91.721 | 85.2 | 91 | 83.7 | 87.1 | 75
| 1 | 6 | 91 | 84.18 | 72.95 | 79.9564 | 83.6583 | 80.2 | 89.9 | 89 | 81 | 79 | 83.4 | 89 | 90 | 85.7 | < 100 | 88.2 | 79 | 71.4387 | 91 | 84 | 81 | 62.55 | 90.9696 | 85 | 85.5857 | 88.3 | 79.203 | 88.2 | 89 | 83.6 | 85.8
| Mean |  | 90.1667 | 84.7467 | 75.1383 | 80.3966 | 84.2359 | 83.9667 | 88.5167 | 88.3333 | 84.3333 | 78.5 | 82.4167 | 90.2167 | 88.8 | 85.2167 | 104.113 | 88.4333 | 79.1667 | 72.8007 | 88 | 85.3833 | 84.5 | 62.005 | 88.3614 | 84.1667 | 89.5867 | 88.55 | 84.1427 | 87.1333 | 88.3333 | 84.75 | 86.25
| Median |  | 90 | 84.655 | 75.665 | 80.5157 | 84.4404 | 83.4 | 88.85 | 89 | 84 | 78.5 | 82.8 | 90.55 | 89.35 | 85.4 | 104.113 | 88.5 | 79 | 72.9409 | 88 | 85.05 | 84 | 62.11 | 89.5436 | 85 | 90.8013 | 88.6 | 83.1475 | 87.35 | 89 | 84.15 | 85.95
| Std Dev. |  | 0.75277 | 0.63999 | 1.5077 | 0.54531 | 1.2841 | 2.6181 | 1.6987 | 1.0328 | 2.6583 | 0.54772 | 1.2416 | 0.94534 | 1.228 | 0.7679 |  | 0.84301 | 0.98319 | 1.5904 | 2.3664 | 2.3008 | 2.5884 | 0.50035 | 3.7905 | 2.0412 | 2.6228 | 1.3368 | 5.9374 | 1.0424 | 2.3381 | 1.7841 | 1.1005
| Rel.Std.Dev. |  | 0.8349% | 0.7552% | 2.007% | 0.6783% | 1.524% | 3.118% | 1.919% | 1.169% | 3.152% | 0.6977% | 1.507% | 1.048% | 1.383% | 0.9011% |  | 0.9533% | 1.242% | 2.185% | 2.689% | 2.695% | 3.063% | 0.807% | 4.29% | 2.425% | 2.928% | 1.51% | 7.056% | 1.196% | 2.647% | 2.105% | 1.276%
| PDM3 |  | 5.438% | -0.9% | -12.14% | -5.987% | -1.497% | -1.812% | 3.508% | 3.294% | -1.383% | -8.205% | -3.625% | 5.496% | 3.84% | -0.3504% | 21.75% | 3.411% | -7.425% | -14.87% | 2.904% | -0.1555% | -1.188% | -27.49% | 3.327% | -1.578% | 4.76% | 3.547% | -1.606% | 1.891% | 3.294% | -0.8961% | 0.8579%
| Z-Score (Absolute) |  | 1.02 | 0 | 1.81 | 0.82 | 0.1 | 0.15 | 0.71 | 0.67 | 0.08 | 1.18 | 0.44 | 1.03 | 0.76 | 0.09 | 4.84 | 0.69 | 1.05 | 2.25 | 0.61 | 0.12 | 0.05 | 4.28 | 0.68 | 0.11 | 0.91 | 0.72 | 0.11 | 0.45 | 0.67 | 0 | 0.28
| Analytical results for B in OREAS 241b (Certified Value 29.9 ppm) | [32 empty] | Certified
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.11 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.23 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 1
| 1 | 1 | 30 | 27 | 37.62 | 28.6109 | 42 | 39 | 30 | 20 | 30 | 30 | 22 | 24 | 83.72 | 20 | 20 | 19 | < 10 | 36 | 22.9 | [12 empty] | 1
| 1 | 2 | 29 | 28 | 39.22 | 30.9426 | 45 | 39 | 30 | 20 | 30 | 30 | 22 | 25 | 78.34 | 20 | 20 | 20 | < 10 | 35 | 23.1 | [12 empty] | 3
| 1 | 3 | 30 | 28 | 37.66 | 28.777 | 42 | 35 | 30 | 20 | 30 | 30 | 26 | 24 | 91.35 | 20 | 20 | 21 | < 10 | 35 | 22.9 | [12 empty] | 16
| 1 | 4 | 29 | 28 | 39.59 | 29.1816 | 40 | 37 | 30 | 20 | 30 | 30 | 26 | 23 | 81.54 | 20 | 20 | 20 | < 10 | 34 | 22.5 | [12 empty] | 29.9367
| 1 | 5 | 29 | 28 | 37.11 | 29.9259 | 42 | 34 | 30 | 20 | 30 | 30 | 22 | 24 | 76.89 | 20 | 20 | 22 | < 10 | 36 | 22 | [12 empty] | 76
| 1 | 6 | 28 | 27 | 38.12 | 27.7879 | 40 | 37 | 30 | 20 | 30 | 30 | 25 | 24 | 89.27 | 20 | 20 | 18 | < 10 | 38 | 23.5
| Mean |  | 29.1667 | 27.6667 | 38.22 | 29.2043 | 41.8333 | 36.8333 | 30 | 20 | 30 | 30 | 23.8333 | 24 | 83.5183 | 20 | 20 | 20 |  | 35.6667 | 22.8167
| Median |  | 29 | 28 | 37.89 | 28.9793 | 42 | 37 | 30 | 20 | 30 | 30 | 23.5 | 24 | 82.63 | 20 | 20 | 20 |  | 35.5 | 22.9
| Std Dev. |  | 0.75277 | 0.5164 | 0.97904 | 1.1031 | 1.8348 | 2.0412 | 0 | 0 | 0 | 0 | 2.0412 | 0.63246 | 5.8159 | 0 | 0 | 1.4142 |  | 1.3663 | 0.51543
| Rel.Std.Dev. |  | 2.581% | 1.866% | 2.562% | 3.777% | 4.386% | 5.542% | 0% | 0% | 0% | 0% | 8.565% | 2.635% | 6.964% | 0% | 0% | 7.071% |  | 3.831% | 2.259%
| PDM3 |  | -2.572% | -7.583% | 27.67% | -2.447% | 39.74% | 23.04% | 0.2113% | -33.19% | 0.2113% | 0.2113% | -20.39% | -19.83% | 179% | -33.19% | -33.19% | -33.19% |  | 19.14% | -23.78%
| Z-Score (Absolute) |  | 0 | 0.16 | 0.95 | 0 | 1.33 | 0.8 | 0.09 | NA | NA | NA | 0.56 | 0.54 | 5.7 | NA | NA | 0.96 | 2.54 | 0.68 | 0.67
Results from laboratories 1.15, 1.16, 1.18, 1.27 and 1.29 were removed due to their 10 ppm reading resolution.
| Analytical results for Ba in OREAS 241b (Certified Value 40.8 ppm) | [32 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 41 | 42.36 | 40.49 | 45.46 | 42.1349 | 40.4 | 49 | 45 | 40 | 41 | 43 | 40 | 40 | 40 | 60.167 | 40 | 29 | 42.5812 | 44 | 33.6 | 42 | 39.63 | 42.506 | 34 | 40 | 39 | 47 | 40 | 39 |  |  | 1
| 1 | 2 | 41 | 42.81 | 41.13 | 45.45 | 42.8564 | 40.5 | 49 | 45 | 41 | 41 | 42 | 40 | 40 | 40 | 61.373 | 40 | 29 | 42.71 | 44 | 35.7 | 41 | 40.11 | 40.545 | 35 | 40 | 41 | 44 | 39 | 39 |  |  | 20
| 1 | 3 | 41 | 42.05 | 40.54 | 45.06 | 42.305 | 40.2 | 49 | 44 | 40 | 41 | 43 | 40 | 40 | 40 | 57.944 | 40 | 29 | 43.0668 | 42 | 35.1 | 41 | 38.48 | 40.086 | 37 | 40 | 41 | 46 | 39 | 40 |  |  | 16
| 1 | 4 | 41 | 42.76 | 40.47 | 45.96 | 41.4028 | 38 | 50 | 45 | 40 | 42 | 43 | 40 | 40 | 40 | 58.661 | 40 | 29 | 42.7874 | 38 | 32.1 | 41 | 38.78 | 37.566 | 34 | 40 | 40 | 46 | 39 | 38 |  |  | 40.7607
| 1 | 5 | 42 | 43.32 | 40.41 | 45.82 | 42.2072 | 38.8 | 50 | 45 | 39 | 42 | 43 | 40 | 40 | 40 | 56.78 | 40 | 28 | 40.4894 | 38 | 36.2 | 42 | 37.91 | 39.63 | 36 | 40 | 40 | 47 | 39 | 39 |  |  | 77
| 1 | 6 | 41 | 42.65 | 41.2 | 45.7 | 41.1684 | 37.9 | 49 | 48 | 39 | 42 | 43 | 40 | 40 | 40 | 59.31 | 40 | 28 | 40.7655 | 40 | 35 | 41 | 37.55 | 42.539 | 37 | 40 | 39 | 46 | 38 | 39
| Mean |  | 41.1667 | 42.6583 | 40.7067 | 45.575 | 42.0125 | 39.3 | 49.3333 | 45.3333 | 39.8333 | 41.5 | 42.8333 | 40 | 40 | 40 | 59.0392 | 40 | 28.6667 | 42.0667 | 41 | 34.6167 | 41.3333 | 38.7433 | 40.4787 | 35.5 | 40 | 40 | 46 | 39 | 39
| Median |  | 41 | 42.705 | 40.515 | 45.58 | 42.1711 | 39.5 | 49 | 45 | 40 | 41.5 | 43 | 40 | 40 | 40 | 58.9855 | 40 | 29 | 42.6456 | 41 | 35.05 | 41 | 38.63 | 40.3155 | 35.5 | 40 | 40 | 46 | 39 | 39
| Std Dev. |  | 0.40825 | 0.43134 | 0.35814 | 0.32185 | 0.62207 | 1.2133 | 0.5164 | 1.3663 | 0.75277 | 0.54772 | 0.40825 | 0 | 0 | 0 | 1.6272 | 0 | 0.5164 | 1.1295 | 2.7568 | 1.5118 | 0.5164 | 0.98419 | 1.8824 | 1.3784 | 0 | 0.89443 | 1.0954 | 0.63246 | 0.63246
| Rel.Std.Dev. |  | 0.9917% | 1.011% | 0.8798% | 0.7062% | 1.481% | 3.087% | 1.047% | 3.014% | 1.89% | 1.32% | 0.9531% | 0% | 0% | 0% | 2.756% | 0% | 1.801% | 2.685% | 6.724% | 4.367% | 1.249% | 2.54% | 4.65% | 3.883% | 0% | 2.236% | 2.381% | 1.622% | 1.622%
| PDM3 |  | 0.9959% | 4.655% | -0.1327% | 11.81% | 3.071% | -3.584% | 21.03% | 11.22% | -2.275% | 1.814% | 5.085% | -1.866% | -1.866% | -1.866% | 44.84% | -1.866% | -29.67% | 3.204% | 0.587% | -15.07% | 1.405% | -4.949% | -0.692% | -12.91% | -1.866% | -1.866% | 12.85% | -4.32% | -4.32%
| Z-Score (Absolute) |  | 0.31 | 0.99 | 0.1 | 2.32 | 0.7 | 0.54 | 4.04 | 2.21 | 0.29 | 0.47 | 1.07 | 0.22 | 0.22 | 0.22 | 8.46 | 0.22 | 5.39 | 0.72 | 0.24 | 2.67 | 0.39 | 0.79 | 0 | 2.27 | 0.22 | 0.22 | 2.52 | 0.67 | 0.67
| Analytical results for Be in OREAS 241b (Certified Value 0.25 ppm) | [32 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 2
| 1 | 1 | 0.2 | 0.28 | < 1 | 0.17758 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.29 | < 5 | < 1 | 0.27 | 0.26 | 0.25 | < 10 | 0.26 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.25 | 0.13 | 0.3 | 0.3 | 0.25 | [5 empty] | 1
| 1 | 2 | 0.2 | 0.27 | < 1 | 0.17824 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.29 | < 5 | < 1 | 0.25 | 0.26 | 0.25 | < 10 | 0.26 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.11 | 0.3 | 0.28 | 0.22 | [5 empty] | 21
| 1 | 3 | 0.2 | 0.31 | < 1 | 0.17498 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.28 | < 5 | < 1 | 0.26 | 0.25 | 0.25 | < 10 | 0.26 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.12 | 0.4 | 0.29 | 0.21 | [5 empty] | 16
| 1 | 4 | 0.2 | 0.3 | < 1 | 0.18264 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.29 | < 5 | < 1 | 0.25 | 0.26 | 0.25 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.2 | 0.24 | 0.13 | 0.3 | 0.3 | 0.19 | [5 empty] | 0.25146
| 1 | 5 | 0.3 | 0.26 | < 1 | 0.18068 | < 0.5 | 0.3 | < 0.5 | < 0.5 | 0.26 | < 5 | < 1 | 0.26 | 0.26 | 0.25 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.3 | 0.24 | 0.12 | 0.4 | 0.28 | 0.21 | [5 empty] | 78
| 1 | 6 | 0.2 | 0.29 | < 1 | 0.17541 | < 0.5 | 0.2 | < 0.5 | < 0.5 | 0.29 | < 5 | < 1 | 0.25 | 0.26 | 0.25 | < 10 | 0.25 | < 0.5 | < 1 | < 0.5 | < 0.5 | 0.3 | 0.24 | 0.1 | 0.3 | 0.29 | 0.19
| Mean |  | 0.21667 | 0.285 |  | 0.17826 |  | 0.28333 |  |  | 0.28333 |  |  | 0.25667 | 0.25833 | 0.25 |  | 0.255 |  |  |  |  | 0.23333 | 0.24167 | 0.11833 | 0.33333 | 0.29 | 0.21167
| Median |  | 0.2 | 0.285 |  | 0.17791 |  | 0.3 |  |  | 0.29 |  |  | 0.255 | 0.26 | 0.25 |  | 0.255 |  |  |  |  | 0.2 | 0.24 | 0.12 | 0.3 | 0.29 | 0.21
| Std Dev. |  | 0.040825 | 0.018708 |  | 0.0029808 |  | 0.040825 |  |  | 0.012111 |  |  | 0.008165 | 0.0040825 | 0 |  | 0.0054772 |  |  |  |  | 0.05164 | 0.0040825 | 0.01169 | 0.05164 | 0.0089443 | 0.022286
| Rel.Std.Dev. |  | 18.84% | 6.564% |  | 1.672% |  | 14.41% |  |  | 4.274% |  |  | 3.181% | 1.58% | 0% |  | 2.148% |  |  |  |  | 22.13% | 1.689% | 9.879% | 15.49% | 3.084% | 10.53%
| PDM3 |  | -13.84% | 13.34% |  | -29.11% |  | 12.68% |  |  | 12.68% |  |  | 2.071% | 2.734% | -0.5802% |  | 1.408% |  |  |  |  | -7.208% | -3.894% | -52.94% | 32.56% | 15.33% | -15.82%
| Z-Score (Absolute) |  | NA | 1.05 | 8.01 | 2.4 | 0.08 | NA | 0.08 | 0.08 | 1 | 72.74 | 8.01 | 0.13 | 0.19 | 0.08 | 153.66 | 0.08 | 0.08 | 8.01 | 0.08 | 0.08 | NA | 0.35 | 4.34 | NA | 1.21 | 1.32
Results from laboratories 1.01, 1.07, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Bi in OREAS 241b (Certified Value 0.063 ppm) | [32 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.05 | 0.07 | < 1 | 0.049213 | 0.06 | < 1 | 0.06 | 0.07 | < 2 | < 0.1 | 0.06 | 0.06 | 0.07 | < 10 | 0.07 | < 5 | 0.06 | 0.05 | 0.07 | < 0.1 | 0.06926 | 0.07 | 0.048 | 0.07 | 0.07 | 0.07 | 0.05 |  |  |  |  | 1
| 1 | 2 | 0.04 | 0.07 | < 1 | 0.048106 | 0.07 | 4 | 0.08 | 0.07 | < 2 | < 0.1 | 0.06 | 0.07 | 0.07 | < 10 | 0.07 | < 5 | 0.06 | 0.05 | 0.07 | < 0.1 | 0.06932 | 0.06 | 0.048 | 0.07 | 0.07 | 0.07 | 0.06 |  |  |  |  | 22
| 1 | 3 | 0.05 | 0.07 | < 1 | 0.044634 | 0.07 | 5 | 0.06 | 0.06 | < 2 | < 0.1 | 0.06 | 0.06 | 0.07 | < 10 | 0.07 | < 5 | 0.07 | 0.06 | 0.07 | < 0.1 | 0.07098 | 0.07 | 0.045 | 0.07 | 0.07 | 0.07 | 0.05 |  |  |  |  | 16
| 1 | 4 | 0.05 | 0.07 | < 1 | 0.044397 | 0.06 | 4 | 0.06 | 0.06 | < 2 | < 0.1 | 0.06 | 0.06 | 0.07 | < 10 | 0.06 | < 5 | 0.06 | 0.05 | 0.07 | < 0.1 | 0.0714 | 0.06 | 0.05 | 0.07 | 0.07 | 0.07 | 0.05 |  |  |  |  | 0.062926
| 1 | 5 | 0.05 | 0.07 | < 1 | 0.045344 | 0.06 | 1 | 0.06 | 0.06 | < 2 | < 0.1 | 0.06 | 0.07 | 0.07 | < 10 | 0.06 | < 5 | 0.06 | 0.05 | 0.06 | < 0.1 | 0.0679 | 0.06 | 0.048 | 0.07 | 0.07 | 0.07 | 0.06 |  |  |  |  | 79
| 1 | 6 | 0.05 | 0.07 | < 1 | 0.042079 | 0.06 | 4 | 0.06 | 0.06 | < 2 | < 0.1 | 0.06 | 0.07 | 0.07 | < 10 | 0.07 | < 5 | 0.06 | 0.05 | 0.07 | < 0.1 | 0.0679 | 0.07 | 0.039 | 0.07 | 0.07 | 0.06 | 0.05
| Mean |  | 0.048333 | 0.07 |  | 0.045629 | 0.063333 | 3.6 | 0.063333 | 0.063333 |  |  | 0.06 | 0.065 | 0.07 |  | 0.066667 |  | 0.061667 | 0.051667 | 0.068333 |  | 0.06946 | 0.065 | 0.046333 | 0.07 | 0.07 | 0.068333 | 0.053333
| Median |  | 0.05 | 0.07 |  | 0.044989 | 0.06 | 4 | 0.06 | 0.06 |  |  | 0.06 | 0.065 | 0.07 |  | 0.07 |  | 0.06 | 0.05 | 0.07 |  | 0.06929 | 0.065 | 0.048 | 0.07 | 0.07 | 0.07 | 0.05
| Std Dev. |  | 0.0040825 | 0 |  | 0.0026144 | 0.005164 | 1.5166 | 0.008165 | 0.005164 |  |  | 0 | 0.0054772 | 0 |  | 0.005164 |  | 0.0040825 | 0.0040825 | 0.0040825 |  | 0.0014833 | 0.0054772 | 0.0039328 | 0 | 0 | 0.0040825 | 0.005164
| Rel.Std.Dev. |  | 8.447% | 0% |  | 5.73% | 8.154% | 42.13% | 12.89% | 8.154% |  |  | 0% | 8.427% | 0% |  | 7.746% |  | 6.62% | 7.902% | 5.974% |  | 2.135% | 8.427% | 8.488% | 0% | 0% | 5.974% | 9.682%
| PDM3 |  | -23.19% | 11.24% |  | -27.49% | 0.6474% | 5621% | 0.6474% | 0.6474% |  |  | -4.65% | 3.296% | 11.24% |  | 5.945% |  | -2.001% | -17.89% | 8.593% |  | 10.38% | 3.296% | -26.37% | 11.24% | 11.24% | 8.593% | -15.24%
| Z-Score (Absolute) |  | 2.25 | 0.67 | 58.66 | 2.61 | 0.22 | NA | 0.22 | 0.22 | 126.1 | 2.02 | 0.67 | 0 | 0.67 | 665.54 | 0.22 | 328.39 | 0.45 | 1.8 | 0.45 | 2.02 | 0.6 | 0 | 2.52 | 0.67 | 0.67 | 0.45 | 1.57
Results from laboratory 1.08 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 241b (Certified Value 2.31 wt.%) | [32 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 2.38 | 2.322 | 2.3824 | 2.7736 | 2.3907 | 2.86 | 2.601 | 2.99 | 2.58 | 2.43 | 2.08 | 2.35 | 2.31 | 2.26 | 1.9269 | 2.32 | 2.24 | 2.5523 | 2.68 | 2.09 | 2.26 | 2.1039 | 2.5653 | 1.6245 | 2.29 | 1.91 | 2.2 | 2.35 | 1.93 |  |  | 1
| 1 | 2 | 2.33 | 2.328 | 2.408 | 2.7882 | 2.4877 | 3.05 | 2.637 | 2.94 | 2.65 | 2.44 | 2.09 | 2.26 | 2.3 | 2.25 | 1.9356 | 2.36 | 2.26 | 2.7493 | 2.58 | 2.28 | 2.24 | 2.1421 | 2.5221 | 1.655 | 2.3 | 2.04 | 2.19 | 2.25 | 1.88 |  |  | #N/A
| 1 | 3 | 2.3 | 2.321 | 2.4192 | 2.781 | 2.4373 | 3.13 | 2.616 | 2.91 | 2.56 | 2.48 | 2.09 | 2.3 | 2.25 | 2.24 | 1.8222 | 2.35 | 2.25 | 2.7441 | 2.52 | 2.16 | 2.28 | 2.1097 | 2.4818 | 1.693 | 2.3 | 2.04 | 2.14 | 2.24 | 1.93 |  |  | 16
| 1 | 4 | 2.32 | 2.345 | 2.3608 | 2.7902 | 2.4392 | 2.97 | 2.652 | 2.92 | 2.54 | 2.49 | 2.1 | 2.27 | 2.28 | 2.27 | 1.8765 | 2.32 | 2.24 | 2.6295 | 2.6 | 2.05 | 2.27 | 1.9606 | 2.4069 | 1.602 | 2.26 | 2.04 | 2.19 | 2.27 | 1.84 |  |  | 2.3084
| 1 | 5 | 2.35 | 2.33 | 2.3845 | 2.8023 | 2.4846 | 2.93 | 2.694 | 2.99 | 2.48 | 2.44 | 2.13 | 2.28 | 2.29 | 2.27 | 1.7764 | 2.26 | 2.15 | 2.5357 | 2.67 | 2.24 | 2.3 | 1.7674 | 2.3956 | 1.625 | 2.3 | 2.06 | 2.13 | 2.28 | 1.92 |  |  | 80
| 1 | 6 | 2.25 | 2.305 | 2.4581 | 2.7802 | 2.3853 | 2.79 | 2.709 | 3.07 | 2.48 | 2.46 | 2.13 | 2.3 | 2.3 | 2.23 | 1.8957 | 2.28 | 2.12 | 2.5019 | 2.52 | 2.19 | 2.25 | 1.915 | 2.4573 | 1.702 | 2.27 | 1.82 | 2.12 | 2.25 | 1.96
| Mean |  | 2.3217 | 2.3252 | 2.4022 | 2.7859 | 2.4375 | 2.955 | 2.6515 | 2.97 | 2.5483 | 2.4567 | 2.1033 | 2.2933 | 2.2883 | 2.2533 | 1.8722 | 2.315 | 2.21 | 2.6188 | 2.595 | 2.1683 | 2.2667 | 1.9998 | 2.4715 | 1.6502 | 2.2867 | 1.985 | 2.1617 | 2.2733 | 1.91
| Median |  | 2.325 | 2.325 | 2.3963 | 2.7846 | 2.4383 | 2.95 | 2.6445 | 2.965 | 2.55 | 2.45 | 2.095 | 2.29 | 2.295 | 2.255 | 1.8861 | 2.32 | 2.24 | 2.5909 | 2.59 | 2.175 | 2.265 | 2.0322 | 2.4695 | 1.64 | 2.295 | 2.04 | 2.165 | 2.26 | 1.925
| Std Dev. |  | 0.04446 | 0.013106 | 0.03425 | 0.010005 | 0.043961 | 0.1239 | 0.04274 | 0.059666 | 0.064627 | 0.024221 | 0.021602 | 0.032042 | 0.02137 | 0.01633 | 0.062061 | 0.038859 | 0.05933 | 0.10758 | 0.069785 | 0.087502 | 0.021602 | 0.14555 | 0.065727 | 0.04039 | 0.017512 | 0.097519 | 0.035449 | 0.040332 | 0.042895
| Rel.Std.Dev. |  | 1.915% | 0.5637% | 1.426% | 0.3591% | 1.804% | 4.193% | 1.612% | 2.009% | 2.536% | 0.9859% | 1.027% | 1.397% | 0.9339% | 0.7247% | 3.315% | 1.679% | 2.685% | 4.108% | 2.689% | 4.035% | 0.9531% | 7.278% | 2.659% | 2.448% | 0.7658% | 4.913% | 1.64% | 1.774% | 2.246%
| PDM3 |  | 0.574% | 0.7256% | 4.061% | 20.69% | 5.59% | 28.01% | 14.86% | 28.66% | 10.39% | 6.422% | -8.884% | -0.6534% | -0.87% | -2.386% | -18.9% | 0.2852% | -4.263% | 13.45% | 12.41% | -6.068% | -1.809% | -13.37% | 7.065% | -28.51% | -0.9422% | -14.01% | -6.357% | -1.52% | -17.26%
| Z-Score (Absolute) |  | 0.12 | 0.13 | 0.45 | 2.03 | 0.59 | 2.73 | 1.48 | 2.79 | 1.05 | 0.67 | 0.78 | 0 | 0.02 | 0.17 | 1.74 | 0.09 | 0.34 | 1.34 | 1.25 | 0.52 | 0.11 | 1.21 | 0.74 | 2.65 | 0.03 | 1.27 | 0.54 | 0.08 | 1.58
| Analytical results for Cd in OREAS 241b (Certified Value 0.78 ppm) | [32 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.84 | 0.76 | 1.02 | 0.55939 | 0.80456 | 0.77 | 0.88 | 0.82 | 0.76 | 0.9 | 0.8 | 0.78 | 0.79 | 0.76 | < 10 | 0.83 | < 1 | < 1 | 0.74 | 0.67 | 0.79 | 1.93 | 0.5 | 0.72706 | 0.8 | 0.66 | 0.84 | 0.78 | 0.77 |  |  | 1
| 1 | 2 | 0.88 | 0.766 | 1.02 | 0.56073 | 0.80615 | 0.84 | 0.85 | 0.88 | 0.74 | 0.9 | 0.8 | 0.79 | 0.76 | 0.74 | < 10 | 0.84 | < 1 | < 1 | 0.78 | 0.68 | 0.77 | 1.9 | 0.5 | 0.73586 | 0.81 | 0.7 | 0.78 | 0.74 | 0.79 |  |  | 23
| 1 | 3 | 0.82 | 0.782 | 1.01 | 0.56638 | 0.76782 | 0.77 | 0.86 | 0.82 | 0.74 | 0.9 | 0.8 | 0.8 | 0.76 | 0.77 | < 10 | 0.85 | < 1 | < 1 | 0.73 | 0.68 | 0.75 | 1.79 | 0.5 | 0.76173 | 0.78 | 0.66 | 0.82 | 0.73 | 0.8 |  |  | 16
| 1 | 4 | 0.88 | 0.791 | < 1 | 0.57302 | 0.78877 | 0.8 | 0.89 | 0.84 | 0.73 | 0.9 | 0.8 | 0.78 | 0.78 | 0.76 | < 10 | 0.81 | < 1 | < 1 | 0.71 | 0.66 | 0.77 | 1.79 | 0.5 | 0.74168 | 0.78 | 0.75 | 0.83 | 0.73 | 0.81 |  |  | 0.77943
| 1 | 5 | 0.83 | 0.817 | 1.12 | 0.55901 | 0.80549 | 0.84 | 0.86 | 0.85 | 0.75 | 0.9 | 0.8 | 0.81 | 0.78 | 0.73 | < 10 | 0.84 | < 1 | < 1 | 0.7 | 0.67 | 0.8 | 1.84 | 0.5 | 0.83835 | 0.78 | 0.71 | 0.81 | 0.74 | 0.82 |  |  | 81
| 1 | 6 | 0.84 | 0.775 | 1.02 | 0.55939 | 0.78859 | 0.78 | 0.92 | 0.86 | 0.74 | 0.9 | 0.8 | 0.77 | 0.8 | 0.75 | < 10 | 0.84 | < 1 | < 1 | 0.77 | 0.66 | 0.74 | 1.86 | 1 | 0.71269 | 0.79 | 0.71 | 0.79 | 0.73 | 0.77
| Mean |  | 0.84833 | 0.78183 | 1.038 | 0.56299 | 0.79356 | 0.8 | 0.87667 | 0.845 | 0.74333 | 0.9 | 0.8 | 0.78833 | 0.77833 | 0.75167 |  | 0.835 |  |  | 0.73833 | 0.67 | 0.77 | 1.8517 | 0.58333 | 0.7529 | 0.79 | 0.69833 | 0.81167 | 0.74167 | 0.79333
| Median |  | 0.84 | 0.7785 | 1.02 | 0.56006 | 0.79667 | 0.79 | 0.87 | 0.845 | 0.74 | 0.9 | 0.8 | 0.785 | 0.78 | 0.755 |  | 0.84 |  |  | 0.735 | 0.67 | 0.77 | 1.85 | 0.5 | 0.73877 | 0.785 | 0.705 | 0.815 | 0.735 | 0.795
| Std Dev. |  | 0.025626 | 0.020469 | 0.046043 | 0.0056383 | 0.015047 | 0.032863 | 0.02582 | 0.023452 | 0.010328 | 0 | 1.2162e-16 | 0.01472 | 0.016021 | 0.01472 |  | 0.013784 |  |  | 0.031885 | 0.0089443 | 0.022804 | 0.057067 | 0.20412 | 0.044901 | 0.012649 | 0.034303 | 0.023166 | 0.019408 | 0.020656
| Rel.Std.Dev. |  | 3.021% | 2.618% | 4.436% | 1.001% | 1.896% | 4.108% | 2.945% | 2.775% | 1.389% | 0% | 1.52e-14% | 1.867% | 2.058% | 1.958% |  | 1.651% |  |  | 4.319% | 1.335% | 2.961% | 3.082% | 34.99% | 5.964% | 1.601% | 4.912% | 2.854% | 2.617% | 2.604%
| PDM3 |  | 8.841% | 0.3088% | 33.17% | -27.77% | 1.814% | 2.64% | 12.48% | 8.413% | -4.631% | 15.47% | 2.64% | 1.143% | -0.1403% | -3.562% |  | 7.13% |  |  | -5.272% | -14.04% | -1.209% | 137.6% | -25.16% | -3.404% | 1.357% | -10.4% | 4.136% | -4.845% | 1.784%
| Z-Score (Absolute) |  | 0.95 | 0.05 | 2.44 | 3.32 | 0.13 | 0.22 | 1.37 | 0.9 | 0.62 | NA | NA | 0.05 | 0.1 | 0.5 | 63.04 | 0.75 | 4.26 | 4.26 | 0.7 | 1.72 | 0.23 | 15.95 | NA | 0.48 | 0.07 | 1.3 | 0.4 | 0.65 | 0.12
Results from laboratories 1.12, 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 241b (Certified Value 11.2 ppm) | [32 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [8 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [8 empty] | 2
| 1 | 1 | 10.7 | 11.439 | 11.67 | 9.0095 | 11.4143 | 11.5 | 2.26 | 13.2 | 11.273 | 12 | 10.75 | 11.2 | 10.65 | 11.15 | < 10 | 12.06 | 12.14 | 10.5 | 11.95 | 11.7 | 11.68 | 11.95 | 11.3 | [8 empty] | 1
| 1 | 2 | 10.54 | 11.363 | 11.078 | 9.221 | 11.4199 | 12.1 | 2.01 | 13.1 | 11.469 | 12 | 10.7 | 10.85 | 10.8 | 11.35 | < 10 | 12.09 | 11.41 | 10.5 | 11.95 | 11.9 | 11.17 | 11.58 | 11.3 | [8 empty] | 24
| 1 | 3 | 10.51 | 11.462 | 10.445 | 9.0195 | 11.3635 | 11.7 | 1.67 | 13.1 | 11.282 | 12 | 10.8 | 11.05 | 10.75 | 11.2 | < 10 | 10.93 | 12.19 | 10.5 | 11.8 | 12.2 | 11.26 | 11.48 | 11.6 | [8 empty] | 16
| 1 | 4 | 10.61 | 11.557 | 12.817 | 9.1914 | 11.436 | 11.5 | 1.74 | 13 | 11.034 | 12 | 10.7 | 11.35 | 10.8 | 11.05 | < 10 | 10.15 | 11.46 | 10 | 11.6 | 12.1 | 11.22 | 11.63 | 11.1 | [8 empty] | 11.2415
| 1 | 5 | 10.8 | 11.52 | 11.203 | 9.1425 | 11.5938 | 11.4 | 1.41 | 13.4 | 10.768 | 12 | 10.75 | 11.4 | 10.6 | 10.85 | < 10 | 11.69 | 11.68 | 11 | 11.9 | 11.9 | 11.41 | 11.47 | 11.4 | [8 empty] | 82
| 1 | 6 | 10.66 | 11.292 | 11.181 | 9.1838 | 11.1731 | 11.2 | 1.8 | 13.6 | 10.821 | 12 | 10.7 | 11.3 | 10.65 | 11 | < 10 | 10.49 | 11.46 | 10.5 | 11.8 | 11.8 | 10.97 | 11.45 | 11.2
| Mean |  | 10.6367 | 11.4388 | 11.399 | 9.128 | 11.4001 | 11.5667 | 1.815 | 13.2333 | 11.1078 | 12 | 10.7333 | 11.1917 | 10.7083 | 11.1 |  | 11.235 | 11.7233 | 10.5 | 11.8333 | 11.9333 | 11.285 | 11.5933 | 11.3167
| Median |  | 10.635 | 11.4505 | 11.192 | 9.1632 | 11.4171 | 11.5 | 1.77 | 13.15 | 11.1535 | 12 | 10.725 | 11.25 | 10.7 | 11.1 |  | 11.31 | 11.57 | 10.5 | 11.85 | 11.9 | 11.24 | 11.53 | 11.3
| Std Dev. |  | 0.10708 | 0.098386 | 0.79791 | 0.091447 | 0.13583 | 0.30768 | 0.29222 | 0.22509 | 0.27974 | 0 | 0.040825 | 0.20837 | 0.08612 | 0.17321 |  | 0.82993 | 0.355 | 0.31623 | 0.13292 | 0.18619 | 0.2404 | 0.18833 | 0.17224
| Rel.Std.Dev. |  | 1.007% | 0.8601% | 7% | 1.002% | 1.191% | 2.66% | 16.1% | 1.701% | 2.518% | 0% | 0.3804% | 1.862% | 0.8042% | 1.56% |  | 7.387% | 3.028% | 3.012% | 1.123% | 1.56% | 2.13% | 1.624% | 1.522%
| PDM3 |  | -5.381% | 1.755% | 1.401% | -18.8% | 1.411% | 2.892% | -83.85% | 17.72% | -1.189% | 6.747% | -4.521% | -0.4435% | -4.743% | -1.259% |  | -0.05807% | 4.286% | -6.596% | 5.264% | 6.154% | 0.3867% | 3.13% | 0.6684%
| Z-Score (Absolute) |  | 1.06 | 0.3 | 0.24 | 3.61 | 0.24 | 0.52 | 15.99 | 3.34 | 0.26 | NA | 0.89 | 0.12 | 0.93 | 0.27 | 10.6 | 0.04 | 0.78 | 1.29 | 0.97 | 1.14 | 0.04 | 0.56 | 0.1
Results from laboratory 1.13 were removed due to their 1 ppm reading resolution.
| Analytical results for Co in OREAS 241b (Certified Value 28.7 ppm) | [32 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 27.8 | 28.96 | 26.12 | 21.2152 | 29.4429 | 28.9 | 28.8 | 31.8 | 30.1 | 27 | 28.6 | 29.1 | 28.5 | 28.2 | 26.026 | 29 | 27 | 31.6378 | 30.7 | 28.6 | 30 | 26.63 | 26 | 30.6378 | 28.1 | 32.4 | 31.1 | 29.5 | 28.8 |  |  | 1
| 1 | 2 | 27.7 | 28.57 | 26.95 | 21.4579 | 30.1603 | 30.7 | 29.6 | 31.4 | 30.8 | 27 | 28.1 | 29.5 | 27.3 | 28.5 | 25.089 | 29.3 | 27 | 31.3649 | 30.2 | 29.2 | 29 | 26.7 | 26 | 31.1453 | 28.5 | 32.8 | 29.1 | 28.9 | 29 |  |  | 25
| 1 | 3 | 27.9 | 28.32 | 26.49 | 21.2358 | 29.6199 | 30.8 | 29.1 | 32.6 | 29.8 | 27 | 28.9 | 29.5 | 27.6 | 28.4 | 24.564 | 28.9 | 27 | 31.8644 | 30.3 | 29.5 | 30 | 25.08 | 27 | 30.9674 | 28.3 | 33.8 | 30.5 | 28.1 | 29.5 |  |  | 16
| 1 | 4 | 28.1 | 28.21 | 25.35 | 21.8577 | 29.9904 | 29.5 | 29.2 | 31.1 | 29.9 | 27 | 28.3 | 29.3 | 28.4 | 28.5 | 27.545 | 28.7 | 27 | 32.4865 | 29.3 | 27.7 | 29 | 23.67 | 26 | 31.3405 | 27.4 | 33 | 30 | 28.5 | 28.6 |  |  | 28.7241
| 1 | 5 | 28.3 | 28.25 | 26.86 | 21.3414 | 29.9418 | 29.6 | 28.6 | 31.7 | 29.3 | 27 | 28.7 | 29.9 | 28.5 | 28.3 | 25.398 | 28.4 | 27 | 31.0789 | 29.1 | 29.1 | 30 | 27.08 | 26 | 30.3126 | 29.2 | 33.1 | 31 | 28.2 | 29.4 |  |  | 83
| 1 | 6 | 28 | 28.11 | 25.56 | 21.2907 | 29.2907 | 28.4 | 29.5 | 32.1 | 28.9 | 27 | 29 | 30.1 | 28.6 | 28 | 26.123 | 28.7 | 27 | 30.684 | 30.7 | 28.6 | 29 | 27.68 | 27 | 29.7402 | 28 | 32.8 | 29.8 | 28.2 | 28.8
| Mean |  | 27.9667 | 28.4033 | 26.2217 | 21.3998 | 29.741 | 29.65 | 29.1333 | 31.7833 | 29.8 | 27 | 28.6 | 29.5667 | 28.15 | 28.3167 | 25.7908 | 28.8333 | 27 | 31.5194 | 30.05 | 28.7833 | 29.5 | 26.14 | 26.3333 | 30.6906 | 28.25 | 32.9833 | 30.25 | 28.5667 | 29.0167
| Median |  | 27.95 | 28.285 | 26.305 | 21.316 | 29.7808 | 29.55 | 29.15 | 31.75 | 29.85 | 27 | 28.65 | 29.5 | 28.45 | 28.35 | 25.712 | 28.8 | 27 | 31.5013 | 30.25 | 28.85 | 29.5 | 26.665 | 26 | 30.8026 | 28.2 | 32.9 | 30.25 | 28.35 | 28.9
| Std Dev. |  | 0.21602 | 0.31354 | 0.66602 | 0.24055 | 0.34195 | 0.95656 | 0.38816 | 0.52694 | 0.65727 | 0 | 0.34641 | 0.37238 | 0.55408 | 0.19408 | 1.0387 | 0.30768 | 0 | 0.62996 | 0.6921 | 0.63692 | 0.54772 | 1.4855 | 0.5164 | 0.59239 | 0.59582 | 0.46655 | 0.76616 | 0.54283 | 0.36009
| Rel.Std.Dev. |  | 0.7724% | 1.104% | 2.54% | 1.124% | 1.15% | 3.226% | 1.332% | 1.658% | 2.206% | 0% | 1.211% | 1.259% | 1.968% | 0.6854% | 4.027% | 1.067% | 0% | 1.999% | 2.303% | 2.213% | 1.857% | 5.683% | 1.961% | 1.93% | 2.109% | 1.414% | 2.533% | 1.9% | 1.241%
| PDM3 |  | -2.637% | -1.117% | -8.712% | -25.5% | 3.54% | 3.223% | 1.425% | 10.65% | 3.746% | -6.002% | -0.4321% | 2.933% | -1.999% | -1.418% | -10.21% | 0.3803% | -6.002% | 9.732% | 4.616% | 0.2062% | 2.701% | -8.996% | -8.323% | 6.846% | -1.651% | 14.83% | 5.312% | -0.5481% | 1.019%
| Z-Score (Absolute) |  | 0.58 | 0.27 | 1.8 | 5.2 | 0.67 | 0.61 | 0.25 | 2.11 | 0.72 | 1.26 | 0.13 | 0.55 | 0.45 | 0.33 | 2.11 | 0.04 | 1.26 | 1.93 | 0.89 | 0 | 0.5 | 1.86 | 1.73 | 1.34 | 0.38 | 2.96 | 1.03 | 0.15 | 0.16
| Analytical results for Cr in OREAS 241b (Certified Value 33.8 ppm) | [32 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 2
| 1 | 1 | 36 | 35 | 35.96 | 34.2076 | 37 | 37 | 39 | 36 | 31 | 36 | 33 | 33 | 31 | 31.659 | 33 | 30 | 34.7869 | 36 | 33.7 | 33 | 34.26 | 34.782 | 30 | 33 | 31.6 | 34 | 32 | 31.7 |  |  |  | 1
| 1 | 2 | 36 | 34.3 | 35.68 | 35.2536 | 39 | 39 | 39 | 37 | 32 | 35 | 32 | 33 | 32 | 32.427 | 33 | 30 | 34.621 | 37 | 36.1 | 32 | 33.78 | 33.2475 | 30 | 33 | 33.5 | 34 | 33 | 31.2 |  |  |  | 26
| 1 | 3 | 37 | 33.8 | 36.09 | 34.5316 | 40 | 39 | 40 | 35 | 33 | 37 | 33 | 32 | 32 | 29.865 | 32 | 30 | 35.0645 | 36 | 34.9 | 33 | 33.66 | 33.5823 | 30 | 34 | 33.2 | 35 | 31 | 32.5 |  |  |  | 16
| 1 | 4 | 37 | 33.9 | 34.75 | 34.9854 | 37 | 38 | 38 | 35 | 32 | 36 | 32 | 33 | 32 | 31.397 | 33 | 31 | 35.1946 | 33 | 33 | 35 | 36.08 | 32.0385 | 30 | 33 | 33.3 | 34 | 32 | 31.4 |  |  |  | 33.7898
| 1 | 5 | 38 | 34.1 | 35.37 | 35.0043 | 38 | 40 | 40 | 36 | 32 | 36 | 33 | 33 | 33 | 29.362 | 32 | 29 | 33.7217 | 34 | 35.2 | 33 | 34.96 | 31.4526 | 30 | 33 | 34.7 | 35 | 32 | 33 |  |  |  | 84
| 1 | 6 | 37 | 34.1 | 34.8 | 33.8451 | 36 | 38 | 41 | 39 | 32 | 37 | 33 | 32 | 32 | 32.658 | 32 | 30 | 33.2684 | 35 | 34.6 | 34 | 33.44 | 32.9635 | 30 | 33 | 32 | 34 | 31 | 32.4
| Mean |  | 36.8333 | 34.2 | 35.4417 | 34.6379 | 37.8333 | 38.5 | 39.5 | 36.3333 | 32 | 36.1667 | 32.6667 | 32.6667 | 32 | 31.228 | 32.5 | 30 | 34.4428 | 35.1667 | 34.5833 | 33.3333 | 34.3633 | 33.0111 | 30 | 33.1667 | 33.05 | 34.3333 | 31.8333 | 32.0333
| Median |  | 37 | 34.1 | 35.525 | 34.7585 | 37.5 | 38.5 | 39.5 | 36 | 32 | 36 | 33 | 33 | 32 | 31.528 | 32.5 | 30 | 34.7039 | 35.5 | 34.75 | 33 | 34.02 | 33.1055 | 30 | 33 | 33.25 | 34 | 32 | 32.05
| Std Dev. |  | 0.75277 | 0.42895 | 0.5729 | 0.54021 | 1.472 | 1.0488 | 1.0488 | 1.5055 | 0.63246 | 0.75277 | 0.5164 | 0.5164 | 0.63246 | 1.3444 | 0.54772 | 0.63246 | 0.77473 | 1.472 | 1.1017 | 1.0328 | 1.0001 | 1.1743 | 0 | 0.40825 | 1.1149 | 0.5164 | 0.75277 | 0.70616
| Rel.Std.Dev. |  | 2.044% | 1.254% | 1.616% | 1.56% | 3.891% | 2.724% | 2.655% | 4.144% | 1.976% | 2.081% | 1.581% | 1.581% | 1.976% | 4.305% | 1.685% | 2.108% | 2.249% | 4.186% | 3.186% | 3.098% | 2.91% | 3.557% | 0% | 1.231% | 3.373% | 1.504% | 2.365% | 2.204%
| PDM3 |  | 9.007% | 1.214% | 4.889% | 2.51% | 11.97% | 13.94% | 16.9% | 7.527% | -5.297% | 7.034% | -3.324% | -3.324% | -5.297% | -7.582% | -3.817% | -11.22% | 1.933% | 4.075% | 2.348% | -1.351% | 1.697% | -2.305% | -11.22% | -1.844% | -2.189% | 1.609% | -5.79% | -5.198%
| Z-Score (Absolute) |  | 1.34 | 0.19 | 0.73 | 0.38 | 1.78 | 2.08 | 2.51 | 1.13 | 0.77 | 1.05 | 0.48 | 0.48 | 0.77 | 1.11 | 0.56 | 1.65 | 0.3 | 0.61 | 0.36 | 0.19 | 0.26 | 0.33 | 1.65 | 0.26 | 0.31 | 0.25 | 0.85 | 0.76
| Analytical results for Cs in OREAS 241b (Certified Value 0.89 ppm) | [32 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 3
| 1 | 1 | 1 | 0.91 | 0.68845 | 0.75 | 0.84 | 0.98 | 0.85 | 1 | 0.87 | 0.89 | 0.87 | 0.9 | 0.88 | 1 | 0.91 | 0.78 | 0.95 | 0.92 | 0.92 | [12 empty] | 1
| 1 | 2 | 1.03 | 0.9 | 0.70186 | 0.8 | 0.84 | 0.9 | 0.91 | 1 | 0.87 | 0.86 | 0.86 | 0.9 | 0.89 | 1 | 0.9 | 0.77 | 0.9 | 0.9 | 0.93 | [12 empty] | 27
| 1 | 3 | 0.97 | 0.91 | 0.69623 | 0.78 | 0.83 | 0.95 | 0.86 | 1.1 | 0.89 | 0.88 | 0.89 | 0.91 | 0.9 | 1 | 0.89 | 0.81 | 0.94 | 0.85 | 0.95 | [12 empty] | 16
| 1 | 4 | 1.02 | 0.92 | 0.74293 | 0.74 | 0.83 | 0.94 | 0.88 | 1 | 0.88 | 0.89 | 0.9 | 0.9 | 0.82 | 1 | 0.9 | 0.77 | 0.96 | 0.89 | 0.91 | [12 empty] | 0.8851
| 1 | 5 | 1.03 | 0.94 | 0.70544 | 0.76 | 0.82 | 0.92 | 0.85 | 1.1 | 0.91 | 0.89 | 0.87 | 0.89 | 0.81 | 1 | 0.91 | 0.79 | 0.97 | 0.87 | 0.92 | [12 empty] | 85
| 1 | 6 | 0.99 | 0.91 | 0.67719 | 0.74 | 0.85 | 0.99 | 0.82 | 1.1 | 0.88 | 0.9 | 0.87 | 0.89 | 0.85 | 1 | 0.91 | 0.74 | 0.94 | 0.87 | 0.93
| Mean |  | 1.0067 | 0.915 | 0.70201 | 0.76167 | 0.835 | 0.94667 | 0.86167 | 1.05 | 0.88333 | 0.885 | 0.87667 | 0.89833 | 0.85833 | 1 | 0.90333 | 0.77667 | 0.94333 | 0.88333 | 0.92667
| Median |  | 1.01 | 0.91 | 0.69904 | 0.755 | 0.835 | 0.945 | 0.855 | 1.05 | 0.88 | 0.89 | 0.87 | 0.9 | 0.865 | 1 | 0.905 | 0.775 | 0.945 | 0.88 | 0.925
| Std Dev. |  | 0.024221 | 0.013784 | 0.022449 | 0.024014 | 0.010488 | 0.034448 | 0.030605 | 0.054772 | 0.015055 | 0.013784 | 0.015055 | 0.0075277 | 0.037639 | 0 | 0.008165 | 0.023381 | 0.024221 | 0.025033 | 0.013663
| Rel.Std.Dev. |  | 2.406% | 1.506% | 3.198% | 3.153% | 1.256% | 3.639% | 3.552% | 5.216% | 1.704% | 1.558% | 1.717% | 0.838% | 4.385% | 0% | 0.9039% | 3.01% | 2.568% | 2.834% | 1.474%
| PDM3 |  | 13.73% | 3.378% | -20.69% | -13.95% | -5.661% | 6.955% | -2.648% | 18.63% | -0.2001% | -0.01177% | -0.9533% | 1.495% | -3.025% | 12.98% | 2.06% | -12.25% | 6.579% | -0.2001% | 4.696%
| Z-Score (Absolute) |  | 2.25 | 0.57 | 3.35 | 2.25 | 0.9 | 1.15 | 0.41 | 3.05 | 0.02 | 0.02 | 0.14 | 0.26 | 0.48 | NA | 0.35 | 1.98 | 1.09 | 0.02 | 0.78
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Cu in OREAS 241b (Certified Value 165 ppm) | [32 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 186 | 171.4 | 158.72 | 179.21 | 167.198 | 158 | 170 | 163 | 168 | 158 | 154.9 | 173.5 | 167 | 165.5 | 173.34 | 164 | 173 | 162.7899 | 162 | 156.69 | 170 | 177.41 | 160.3932 | 160 | 169.5316 | 168.5 | 142.224 | 170 | 170 | 165.3 | 165 | 1
| 1 | 2 | 182.8 | 170.27 | 161.29 | 175.01 | 174.4739 | 168 | 173 | 159 | 172 | 161 | 153.1 | 170.5 | 167 | 167 | 175.619 | 164.5 | 174 | 162.6563 | 165 | 158.66 | 166 | 175.71 | 158.958 | 161 | 171.9749 | 169 | 164.369 | 174 | 162.5 | 162.5 | 165 | 28
| 1 | 3 | 184.2 | 169.1 | 157.67 | 171.09 | 168.1318 | 163 | 170 | 165 | 165.6 | 160 | 155.8 | 172 | 166.5 | 164.5 | 164.766 | 163 | 172 | 163.4889 | 164 | 157.5 | 165 | 172.48 | 160.5384 | 164 | 169.7768 | 168.5 | 146.113 | 173 | 168.5 | 161.2 | 170 | 16
| 1 | 4 | 185.7 | 168.45 | 157.65 | 170.58 | 171.4044 | 161 | 170 | 158 | 167.8 | 162 | 153.3 | 170 | 164 | 164 | 172.128 | 163 | 176 | 169.6046 | 162 | 152.34 | 165 | 172.59 | 157.642 | 160 | 168.9342 | 165.5 | 148.219 | 171 | 170 | 163.9 | 165 | 165.3704
| 1 | 5 | 188.3 | 167.92 | 172.13 | 175.23 | 170.8234 | 162 | 171 | 161 | 163.9 | 161 | 154.4 | 170 | 164.5 | 168 | 163.743 | 161.5 | 174 | 162.9575 | 164 | 158.01 | 168 | 171.19 | 157.2464 | 162 | 169.0537 | 169 | 164.874 | 172 | 172.6 | 161.1 | 170 | 86
| 1 | 6 | 189.2 | 166.54 | 155.73 | 174.38 | 166.726 | 157 | 169 | 162 | 162.7 | 162 | 157.8 | 170.5 | 165.5 | 165.5 | 171.719 | 165 | 175 | 160.8125 | 165 | 153.92 | 165 | 170.95 | 155.996 | 165 | 164.2524 | 168.5 | 142.233 | 168 | 165.2 | 159.2 | 165
| Mean |  | 186.0333 | 168.9467 | 160.5317 | 174.25 | 169.7929 | 161.5 | 170.5 | 161.3333 | 166.6667 | 160.6667 | 154.8833 | 171.0833 | 165.75 | 165.75 | 170.2192 | 163.5 | 174 | 163.7183 | 163.6667 | 156.1867 | 166.5 | 173.3883 | 158.4623 | 162 | 168.9206 | 168.1667 | 151.3387 | 171.3333 | 168.1333 | 162.2 | 166.6667
| Median |  | 185.85 | 168.775 | 158.195 | 174.695 | 169.4776 | 161.5 | 170 | 161.5 | 166.7 | 161 | 154.65 | 170.5 | 166 | 165.5 | 171.9235 | 163.5 | 174 | 162.8737 | 164 | 157.095 | 165.5 | 172.535 | 158.3 | 161.5 | 169.2926 | 168.5 | 147.166 | 171.5 | 169.25 | 161.85 | 165
| Std Dev. |  | 2.4122 | 1.7255 | 5.9654 | 3.15 | 2.9823 | 3.937 | 1.3784 | 2.582 | 3.3476 | 1.5055 | 1.7475 | 1.3934 | 1.2942 | 1.5083 | 4.8264 | 1.2649 | 1.4142 | 3.0237 | 1.3663 | 2.5042 | 2.0736 | 2.6002 | 1.8176 | 2.0976 | 2.5395 | 1.3292 | 10.5454 | 2.1602 | 3.6713 | 2.1817 | 2.582
| Rel.Std.Dev. |  | 1.297% | 1.021% | 3.716% | 1.808% | 1.756% | 2.438% | 0.8084% | 1.6% | 2.009% | 0.9371% | 1.128% | 0.8145% | 0.7808% | 0.91% | 2.835% | 0.7736% | 0.8128% | 1.847% | 0.8348% | 1.603% | 1.245% | 1.5% | 1.147% | 1.295% | 1.503% | 0.7904% | 6.968% | 1.261% | 2.184% | 1.345% | 1.549%
| PDM3 |  | 12.49% | 2.163% | -2.926% | 5.369% | 2.674% | -2.34% | 3.102% | -2.441% | 0.7838% | -2.844% | -6.342% | 3.455% | 0.2295% | 0.2295% | 2.932% | -1.131% | 5.218% | -0.9991% | -1.03% | -5.553% | 0.683% | 4.848% | -4.177% | -2.038% | 2.147% | 1.691% | -8.485% | 3.606% | 1.671% | -1.917% | 0.7838%
| Z-Score (Absolute) |  | 3.06 | 0.38 | 0.94 | 1.22 | 0.52 | 0.78 | 0.63 | 0.81 | 0.03 | 0.91 | 1.82 | 0.72 | 0.12 | 0.12 | 0.58 | 0.47 | 1.18 | 0.44 | 0.44 | 1.62 | 0 | 1.08 | 1.26 | 0.71 | 0.38 | 0.26 | 2.38 | 0.76 | 0.26 | 0.67 | 0.03
| Analytical results for Dy in OREAS 241b (Certified Value 2.12 ppm) | [32 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 2.233 | 1.969 | 1.7214 | 2.4 | 2.59 | 1.8 | [25 empty] | 1
| 1 | 2 | 2.221 | 2.04 | 1.765 | 2.5 | 2.47 | 1.8 | [25 empty] | 29
| 1 | 3 | 2.151 | 1.912 | 1.7231 | 2.5 | 2.56 | 1.7 | [25 empty] | 16
| 1 | 4 | 2.239 | 2.015 | 1.7118 | 2.4 | 2.52 | 1.7 | [25 empty] | 2.1191
| 1 | 5 | 2.277 | 2.032 | 1.7709 | 2.4 | 2.58 | 1.8 | [25 empty] | 87
| 1 | 6 | 2.202 | 2.026 | 1.7655 | 2.3 | 2.59 | 1.8
| Mean |  | 2.2205 | 1.999 | 1.743 | 2.4167 | 2.5517 | 1.7667
| Median |  | 2.227 | 2.0205 | 1.744 | 2.4 | 2.57 | 1.8
| Std Dev. |  | 0.042103 | 0.049445 | 0.026843 | 0.075277 | 0.047924 | 0.05164
| Rel.Std.Dev. |  | 1.896% | 2.473% | 1.54% | 3.115% | 1.878% | 2.923%
| PDM3 |  | 4.783% | -5.669% | -17.75% | 14.04% | 20.41% | -16.63%
| Z-Score (Absolute) |  | 0.23 | 0.23 | 0.76 | 0.64 | 0.92 | 0.71
| Analytical results for Er in OREAS 241b (Certified Value 1.22 ppm) | [32 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.28 | 1.116 | 1.0549 | 1.5 | 1.44 | 0.95 | [25 empty] | 1
| 1 | 2 | 1.276 | 1.182 | 1.1076 | 1.4 | 1.45 | 1 | [25 empty] | 30
| 1 | 3 | 1.304 | 1.091 | 1.0475 | 1.4 | 1.48 | 0.95 | [25 empty] | 16
| 1 | 4 | 1.274 | 1.129 | 1.0489 | 1.3 | 1.45 | 0.95 | [25 empty] | 1.2188
| 1 | 5 | 1.308 | 1.159 | 1.0603 | 1.4 | 1.46 | 0.95 | [25 empty] | 88
| 1 | 6 | 1.298 | 1.171 | 1.0628 | 1.4 | 1.48 | 1
| Mean |  | 1.29 | 1.1413 | 1.0636 | 1.4 | 1.46 | 0.96667
| Median |  | 1.289 | 1.144 | 1.0576 | 1.4 | 1.455 | 0.95
| Std Dev. |  | 0.015073 | 0.035138 | 0.022345 | 0.063246 | 0.016733 | 0.02582
| Rel.Std.Dev. |  | 1.168% | 3.079% | 2.101% | 4.518% | 1.146% | 2.671%
| PDM3 |  | 5.841% | -6.357% | -12.73% | 14.87% | 19.79% | -20.69%
| Z-Score (Absolute) |  | 0.3 | 0.3 | 0.61 | 0.74 | 0.98 | 1
| Analytical results for Eu in OREAS 241b (Certified Value 0.48 ppm) | [32 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 0.504 | 0.385 | 0.48655 | 0.6 | 0.56 | 0.4 | [25 empty] | 1
| 1 | 2 | 0.499 | 0.405 | 0.48375 | 0.6 | 0.55 | 0.4 | [25 empty] | 31
| 1 | 3 | 0.5 | 0.383 | 0.47528 | 0.6 | 0.55 | 0.4 | [25 empty] | 16
| 1 | 4 | 0.499 | 0.391 | 0.52453 | 0.6 | 0.56 | 0.3 | [25 empty] | 0.48373
| 1 | 5 | 0.509 | 0.401 | 0.48791 | 0.6 | 0.55 | 0.4 | [25 empty] | 89
| 1 | 6 | 0.497 | 0.403 | 0.48178 | 0.5 | 0.57 | 0.4
| Mean |  | 0.50133 | 0.39467 | 0.48996 | 0.58333 | 0.55667 | 0.38333
| Median |  | 0.4995 | 0.396 | 0.48515 | 0.6 | 0.555 | 0.4
| Std Dev. |  | 0.0044121 | 0.0095847 | 0.017503 | 0.040825 | 0.008165 | 0.040825
| Rel.Std.Dev. |  | 0.8801% | 2.429% | 3.572% | 6.999% | 1.467% | 10.65%
| PDM3 |  | 3.639% | -18.41% | 1.289% | 20.59% | 15.08% | -20.75%
| Z-Score (Absolute) |  | 0.05 | 0.92 | 0.05 | 0.8 | 0.55 | 1.02
| Analytical results for Fe in OREAS 241b (Certified Value 5.57 wt.%) | [32 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 5.77 | 5.552 | 3.7756 | 5.4884 | 5.6139 | 5.51 | 5.728 | 6.14 | 5.62 | 5.35 | 5.32 | 5.6 | 5.43 | 5.43 | 5.8319 | 5.56 | 5.5 | > 5 | 5.87 | 4.98 | 5.54 | 6.39 | 5.9082 | 4.9845 | 5.43 | 5.37 | 5.62 | 5.66 | 5.59 |  |  | 1
| 1 | 2 | 5.73 | 5.518 | 3.8694 | 5.4109 | 5.7863 | 5.83 | 5.847 | 6.09 | 5.77 | 5.26 | 5.3 | 5.45 | 5.42 | 5.48 | 5.9255 | 5.58 | 5.54 | > 5 | 5.92 | 5.13 | 5.49 | 6.82 | 5.9314 | 4.99 | 5.43 | 5.6 | 5.52 | 5.57 | 5.45 |  |  | #N/A
| 1 | 3 | 5.77 | 5.509 | 3.8756 | 5.4236 | 5.6237 | 5.96 | 5.798 | 6.16 | 5.49 | 5.32 | 5.29 | 5.5 | 5.39 | 5.46 | 5.5903 | 5.54 | 5.5 | > 5 | 6.03 | 5.02 | 5.55 | 6.64 | 5.8351 | 5.1545 | 5.46 | 5.56 | 5.43 | 5.57 | 5.54 |  |  | 16
| 1 | 4 | 5.78 | 5.531 | 3.7735 | 5.4092 | 5.6747 | 5.61 | 5.777 | 6.05 | 5.56 | 5.28 | 5.33 | 5.45 | 5.46 | 5.45 | 5.8602 | 5.51 | 5.58 | > 5 | 5.83 | 4.82 | 5.51 | 6.07 | 5.6691 | 4.9455 | 5.37 | 5.56 | 5.56 | 5.58 | 5.45 |  |  | 5.5679
| 1 | 5 | 5.89 | 5.513 | 3.8826 | 5.4643 | 5.6823 | 5.83 | 5.812 | 6.15 | 5.4 | 5.3 | 5.37 | 5.46 | 5.43 | 5.54 | 5.5089 | 5.43 | 5.53 | > 5 | 5.94 | 5.02 | 5.54 | 5.6 | 5.6967 | 5.0195 | 5.42 | 5.45 | 5.56 | 5.56 | 5.51 |  |  | 90
| 1 | 6 | 5.82 | 5.467 | 3.7843 | 5.4255 | 5.5447 | 5.44 | 5.763 | 6.23 | 5.43 | 5.34 | 5.33 | 5.46 | 5.4 | 5.52 | 5.8315 | 5.55 | 5.52 | > 5 | 5.97 | 4.97 | 5.47 | 6.22 | 5.9859 | 5.102 | 5.43 | 5.3 | 5.51 | 5.47 | 5.57
| Mean |  | 5.7933 | 5.515 | 3.8268 | 5.437 | 5.6543 | 5.6967 | 5.7875 | 6.1367 | 5.545 | 5.3083 | 5.3233 | 5.4867 | 5.4217 | 5.48 | 5.7581 | 5.5283 | 5.5283 |  | 5.9267 | 4.99 | 5.5167 | 6.29 | 5.8377 | 5.0327 | 5.4233 | 5.4733 | 5.5333 | 5.5683 | 5.5183
| Median |  | 5.775 | 5.5155 | 3.8268 | 5.4245 | 5.6492 | 5.72 | 5.7875 | 6.145 | 5.525 | 5.31 | 5.325 | 5.46 | 5.425 | 5.47 | 5.8317 | 5.545 | 5.525 |  | 5.93 | 5 | 5.525 | 6.305 | 5.8716 | 5.0047 | 5.43 | 5.505 | 5.54 | 5.57 | 5.525
| Std Dev. |  | 0.055377 | 0.028206 | 0.05402 | 0.032103 | 0.08147 | 0.20646 | 0.041235 | 0.06186 | 0.13693 | 0.034881 | 0.028048 | 0.058538 | 0.024833 | 0.042426 | 0.16707 | 0.053448 | 0.029944 |  | 0.071181 | 0.1008 | 0.032042 | 0.43423 | 0.12963 | 0.079451 | 0.029439 | 0.12028 | 0.06377 | 0.06047 | 0.05947
| Rel.Std.Dev. |  | 0.9559% | 0.5114% | 1.412% | 0.5905% | 1.441% | 3.624% | 0.7125% | 1.008% | 2.469% | 0.6571% | 0.5269% | 1.067% | 0.458% | 0.7742% | 2.901% | 0.9668% | 0.5417% |  | 1.201% | 2.02% | 0.5808% | 6.904% | 2.221% | 1.579% | 0.5428% | 2.198% | 1.152% | 1.086% | 1.078%
| PDM3 |  | 4.049% | -0.95% | -31.27% | -2.351% | 1.551% | 2.313% | 3.944% | 10.22% | -0.4112% | -4.662% | -4.392% | -1.459% | -2.626% | -1.579% | 3.416% | -0.7106% | -0.7106% |  | 6.444% | -10.38% | -0.9201% | 12.97% | 4.846% | -9.613% | -2.596% | -1.698% | -0.6208% | 0.007824% | -0.8902%
| Z-Score (Absolute) |  | 1.57 | 0.05 | 9.84 | 0.5 | 0.76 | 1 | 1.53 | 3.56 | 0.13 | 1.25 | 1.16 | 0.21 | 0.59 | 0.25 | 1.36 | 0.03 | 0.03 | NA | 2.34 | 3.09 | 0.04 | 4.44 | 1.82 | 2.84 | 0.58 | 0.29 | 0.06 | 0.26 | 0.03
| Analytical results for Ga in OREAS 241b (Certified Value 10 ppm) | [32 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 2
| 1 | 1 | 9.7 | 10.6 | 8.811 | 8.7852 | 9.66 | 10.6 | 10.6 | 9.85 | 10 | 10 | 10.15 | 10.2 | 9.41 | 10.2 | 7 | 10.8 | 8.6 | 11.6 | 13.18 | 10 | 9.94 | 11.6 | 10.5 | 7.85 | 8.87 | [6 empty] | 1
| 1 | 2 | 9.4 | 10.6 | 8.8575 | 9.1625 | 10.1 | 10.9 | 10.4 | 10.65 | 10 | 10 | 10.15 | 9.94 | 9.5 | 10.35 | 6 | 10.7 | 9 | 11.4 | 13.32 | 10 | 10.1 | 12.2 | 9.9 | 7.68 | 9.03 | [6 empty] | 32
| 1 | 3 | 9.4 | 10.5 | 8.8746 | 8.9051 | 10.3 | 11.4 | 10.6 | 10.16 | 10 | 10 | 10.2 | 10.05 | 9.47 | 10.3 | 6 | 10.7 | 8.9 | 11.2 | 12.89 | 10 | 9.9 | 12 | 10.2 | 7.39 | 9.15 | [6 empty] | 16
| 1 | 4 | 9.4 | 10.6 | 8.9535 | 8.4854 | 9.58 | 11.1 | 10.4 | 10.05 | 10 | 10 | 10.25 | 10.25 | 9.62 | 10.35 | 6 | 10.3 | 8.3 | 11.6 | 13.1 | 10 | 9.7 | 11.9 | 10.4 | 7.5 | 8.78 | [6 empty] | 10.0383
| 1 | 5 | 9.3 | 10.7 | 8.9273 | 9.023 | 9.7 | 10.9 | 10.5 | 9.74 | 10 | 10 | 10.1 | 10.1 | 9.36 | 9.94 | 7 | 10.3 | 8.8 | 11.7 | 13.08 | 10 | 10 | 12.3 | 10.5 | 7.45 | 9.05 | [6 empty] | 91
| 1 | 6 | 9.2 | 10.5 | 8.9358 | 8.7358 | 9.29 | 11.3 | 10.7 | 9.68 | 10 | 10 | 10 | 10.15 | 9.42 | 10.1 | 6 | 10.5 | 8.5 | 10.9 | 13.06 | 11 | 10.05 | 11.6 | 9.9 | 7.58 | 8.92
| Mean |  | 9.4 | 10.5833 | 8.8933 | 8.8495 | 9.7717 | 11.0333 | 10.5333 | 10.0217 | 10 | 10 | 10.1417 | 10.115 | 9.4633 | 10.2067 | 6.3333 | 10.55 | 8.6833 | 11.4 | 13.105 | 10.1667 | 9.9483 | 11.9333 | 10.2333 | 7.575 | 8.9667
| Median |  | 9.4 | 10.6 | 8.901 | 8.8452 | 9.68 | 11 | 10.55 | 9.95 | 10 | 10 | 10.15 | 10.125 | 9.445 | 10.25 | 6 | 10.6 | 8.7 | 11.5 | 13.09 | 10 | 9.97 | 11.95 | 10.3 | 7.54 | 8.975
| Std Dev. |  | 0.16733 | 0.075277 | 0.054781 | 0.23699 | 0.36696 | 0.29439 | 0.12111 | 0.35785 | 0 | 0 | 0.08612 | 0.11113 | 0.090921 | 0.16269 | 0.5164 | 0.21679 | 0.26394 | 0.30332 | 0.14195 | 0.40825 | 0.14148 | 0.29439 | 0.28048 | 0.16861 | 0.13486
| Rel.Std.Dev. |  | 1.78% | 0.7113% | 0.616% | 2.678% | 3.755% | 2.668% | 1.15% | 3.571% | 0% | 0% | 0.8492% | 1.099% | 0.9608% | 1.594% | 8.154% | 2.055% | 3.04% | 2.661% | 1.083% | 4.016% | 1.422% | 2.467% | 2.741% | 2.226% | 1.504%
| PDM3 |  | -6.359% | 5.429% | -11.41% | -11.84% | -2.656% | 9.912% | 4.931% | -0.1657% | -0.3816% | -0.3816% | 1.03% | 0.764% | -5.728% | 1.677% | -36.91% | 5.097% | -13.5% | 13.56% | 30.55% | 1.279% | -0.8963% | 18.88% | 1.943% | -24.54% | -10.68%
| Z-Score (Absolute) |  | 0.74 | 0.48 | 1.26 | 1.31 | 0.36 | 0.95 | 0.43 | 0.1 | NA | NA | 0.03 | 0 | 0.67 | 0.09 | NA | 0.45 | 1.48 | 1.33 | 3.09 | NA | 0.17 | 1.88 | 0.12 | 2.63 | 1.19
Results from laboratories 1.12, 1.13, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 241b (Certified Value 1.91 ppm) | [32 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 2.057 | 1.583 | 1.9702 | 2.2 | 1.9 | 1.7 | [25 empty] | 1
| 1 | 2 | 2.028 | 1.648 | 1.97 | 2.3 | 1.9 | 1.7 | [25 empty] | 33
| 1 | 3 | 2.043 | 1.533 | 1.9493 | 2.1 | 1.96 | 1.7 | [25 empty] | 16
| 1 | 4 | 2.09 | 1.654 | 1.9194 | 2.1 | 2 | 1.6 | [25 empty] | 1.9099
| 1 | 5 | 2.097 | 1.66 | 1.9455 | 2.2 | 1.98 | 1.7 | [25 empty] | 92
| 1 | 6 | 2.061 | 1.598 | 1.9801 | 2.1 | 1.93 | 1.9
| Mean |  | 2.0627 | 1.6127 | 1.9557 | 2.1667 | 1.945 | 1.7167
| Median |  | 2.059 | 1.623 | 1.9597 | 2.15 | 1.945 | 1.7
| Std Dev. |  | 0.026643 | 0.050278 | 0.022237 | 0.08165 | 0.041833 | 0.098319
| Rel.Std.Dev. |  | 1.292% | 3.118% | 1.137% | 3.768% | 2.151% | 5.727%
| PDM3 |  | 7.999% | -15.56% | 2.4% | 13.44% | 1.838% | -10.12%
| Z-Score (Absolute) |  | 0.46 | 1.39 | 0.02 | 0.89 | 0.02 | 0.96
| Analytical results for Ge in OREAS 241b (Certified Value 0.1 ppm) | [32 empty] | Certified
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 2
| 1 | 1 | < 0.1 | 0.08 | 0.18624 | < 0.2 | < 0.1 | 0.08 | 2 | 0.12 | 0.05 | 0.09 | 0.12 | < 10 | 0.1 | 0.14 | 0.14 | 0.2 | 0.09 | 0.12 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.07 | 0.20184 | < 0.2 | < 0.1 | 0.08 | 2 | 0.12 | 0.05 | 0.11 | 0.13 | < 10 | 0.1 | 0.14 | 0.14 | 0.2 | 0.09 | 0.1 | [13 empty] | 34
| 1 | 3 | < 0.1 | 0.07 | 0.19436 | < 0.2 | < 0.1 | 0.11 | 2.1 | 0.12 | 0.05 | 0.08 | 0.12 | < 10 | 0.1 | 0.14 | 0.12 | 0.2 | 0.09 | 0.11 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.07 | 0.22904 | < 0.2 | < 0.1 | 0.1 | 2 | 0.12 | 0.05 | 0.11 | 0.14 | < 10 | 0.1 | 0.13 | 0.12 | 0.2 | 0.08 | 0.09 | [13 empty] | 0.1019
| 1 | 5 | < 0.1 | 0.08 | 0.21437 | < 0.2 | < 0.1 | 0.1 | 2 | 0.14 | 0.05 | 0.1 | 0.13 | < 10 | 0.1 | 0.13 | 0.1 | 0.2 | 0.08 | 0.1 | [13 empty] | 93
| 1 | 6 | < 0.1 | 0.08 | 0.20261 | < 0.2 | < 0.1 | 0.1 | 2.1 | 0.12 | 0.05 | 0.1 | 0.12 | < 10 | 0.1 | 0.14 | 0.21 | 0.2 | 0.1 | 0.09
| Mean |  |  | 0.075 | 0.20474 |  |  | 0.095 | 2.0333 | 0.12333 | 0.05 | 0.098333 | 0.12667 |  | 0.1 | 0.13667 | 0.13833 | 0.2 | 0.088333 | 0.10167
| Median |  |  | 0.075 | 0.20222 |  |  | 0.1 | 2 | 0.12 | 0.05 | 0.1 | 0.125 |  | 0.1 | 0.14 | 0.13 | 0.2 | 0.09 | 0.1
| Std Dev. |  |  | 0.0054772 | 0.01514 |  |  | 0.012247 | 0.05164 | 0.008165 | 7.6012e-18 | 0.01169 | 0.008165 |  | 1.5202e-17 | 0.005164 | 0.038166 | 3.0405e-17 | 0.0075277 | 0.01169
| Rel.Std.Dev. |  |  | 7.303% | 7.395% |  |  | 12.89% | 2.54% | 6.62% | 1.52e-14% | 11.89% | 6.446% |  | 1.52e-14% | 3.779% | 27.59% | 1.52e-14% | 8.522% | 11.5%
| PDM3 |  |  | -26.4% | 100.9% |  |  | -6.771% | 1895% | 21.03% | -50.93% | -3.5% | 24.3% |  | -1.865% | 34.12% | 35.75% | 96.27% | -13.31% | -0.229%
| Z-Score (Absolute) |  | 1.26 | 0.63 | 2.65 | 0 | 1.26 | 0.13 | NA | 0.59 | 1.26 | 0.04 | 0.67 | 123.9 | NA | 0.93 | 0.97 | NA | 0.3 | 0.04
Results from laboratories 1.13, 1.27 and 1.32 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 241b (Certified Value 0.45 ppm) | [32 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [12 empty] | 2
| 1 | 1 | 0.39 | 0.5 | 0.36988 | 0.4 | 0.38 | 0.47 | 0.61 | 0.4 | 0.46 | 0.49 | 0.49 | 0.47 | 0.48 | < 0.5 | 0.49 | 0.68 | 0.5 | 0.55 | 0.41 | [12 empty] | 1
| 1 | 2 | 0.4 | 0.51 | 0.39013 | 0.4 | 0.39 | 0.44 | 0.63 | 0.4 | 0.45 | 0.53 | 0.47 | 0.46 | 0.41 | < 0.5 | 0.5 | 0.7 | 0.47 | 0.51 | 0.41 | [12 empty] | 18
| 1 | 3 | 0.35 | 0.51 | 0.3733 | 0.5 | 0.41 | 0.47 | 0.6 | 0.4 | 0.46 | 0.48 | 0.48 | 0.43 | 0.46 | < 0.5 | 0.47 | 0.71 | 0.48 | 0.52 | 0.43 | [12 empty] | 16
| 1 | 4 | 0.36 | 0.51 | 0.41508 | 0.5 | 0.68 | 0.46 | 0.61 | 0.4 | 0.47 | 0.49 | 0.48 | 0.47 | 0.46 | < 0.5 | 0.51 | 0.69 | 0.5 | 0.52 | 0.39 | [12 empty] | 0.45457
| 1 | 5 | 0.36 | 0.52 | 0.42512 | 0.4 | 0.4 | 0.49 | 0.61 | 0.4 | 0.45 | 0.49 | 0.46 | 0.44 | 0.42 | < 0.5 | 0.5 | 0.69 | 0.48 | 0.53 | 0.41 | [12 empty] | 94
| 1 | 6 | 0.37 | 0.52 | 0.39274 | 0.4 | 0.41 | 0.49 | 0.62 | 0.4 | 0.47 | 0.5 | 0.45 | 0.43 | 0.42 | < 0.5 | 0.47 | 0.62 | 0.46 | 0.52 | 0.43
| Mean |  | 0.37167 | 0.51167 | 0.39438 | 0.43333 | 0.445 | 0.47 | 0.61333 | 0.4 | 0.46 | 0.49667 | 0.47167 | 0.45 | 0.44167 |  | 0.49 | 0.68167 | 0.48167 | 0.525 | 0.41333
| Median |  | 0.365 | 0.51 | 0.39144 | 0.4 | 0.405 | 0.47 | 0.61 | 0.4 | 0.46 | 0.49 | 0.475 | 0.45 | 0.44 |  | 0.495 | 0.69 | 0.48 | 0.52 | 0.41
| Std Dev. |  | 0.019408 | 0.0075277 | 0.022086 | 0.05164 | 0.11572 | 0.018974 | 0.010328 | 6.0809e-17 | 0.0089443 | 0.017512 | 0.01472 | 0.018974 | 0.028577 |  | 0.016733 | 0.031885 | 0.016021 | 0.013784 | 0.015055
| Rel.Std.Dev. |  | 5.222% | 1.471% | 5.6% | 11.92% | 26% | 4.037% | 1.684% | 1.52e-14% | 1.944% | 3.526% | 3.121% | 4.216% | 6.47% |  | 3.415% | 4.678% | 3.326% | 2.626% | 3.642%
| PDM3 |  | -18.24% | 12.56% | -13.24% | -4.673% | -2.106% | 3.393% | 34.92% | -12.01% | 1.194% | 9.26% | 3.76% | -1.006% | -2.84% |  | 7.793% | 49.96% | 5.96% | 15.49% | -9.072%
| Z-Score (Absolute) |  | 2.34 | 0.99 | 1.8 | NA | 0.59 | 0 | 3.41 | NA | 0.24 | 0.63 | 0.04 | 0.48 | 0.67 | 5.24 | 0.48 | 5.04 | 0.28 | 1.31 | 1.35
Results from laboratories 1.07 and 1.13 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 241b (Indicative Value 0.057 ppm) | [32 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | < 0.01 | 0.02 | < 1 | 0.01 | < 1 | < 0.1 | < 5 | 0.567 | 0.01 | 0.03 | < 0.01 | 0.03 | < 1 | 0.12 | 0.013 | 0.07 | 0.1 | 0.03 | < 0.08 | 0.04 | 0.126 | 0.33 | < 0.01 | 0.024 | 0.56 | [6 empty] | 1
| 1 | 2 | < 0.01 | 0.02 | < 1 | 0.06 | < 1 | < 0.1 | < 5 | 0.561 | 0.02 | 0.02 | < 0.01 | 0.05 | < 1 | 0.11 | 0.013 | 0.08 | 0.14 | 0.04 | < 0.08 | 0.04 | 0.138 | 0.33 | < 0.01 | 0.024 | 0.51 | [6 empty] | 5
| 1 | 3 | < 0.01 | 0.03 | < 1 | 0.09 | < 1 | < 0.1 | < 5 | 0.545 | 0.03 | 0.02 | < 0.01 | 0.03 | < 1 | 0.12 | 0.016 | 0.09 | < 0.1 | 0.04 | < 0.08 | 0.04 | 0.149 | 0.34 | < 0.01 | 0.023 | 0.5 | [6 empty] | 16
| 1 | 4 | < 0.01 | 0.02 | < 1 | 0.09 | < 1 | < 0.1 | < 5 | 0.545 | 0.04 | 0.04 | < 0.01 | 0.02 | < 1 | 0.09 | 0.011 | 0.06 | 0.15 | 0.04 | < 0.08 | 0.04 | 0.136 | 0.33 | < 0.01 | 0.027 | 0.46 | [6 empty] | 0.056923
| 1 | 5 | < 0.01 | 0.02 | < 1 | 0.06 | < 1 | < 0.1 | < 5 | 0.562 | 0.04 | 0.02 | < 0.01 | 0.03 | < 1 | 0.09 | 0.012 | 0.07 | < 0.1 | 0.04 | < 0.08 | 0.04 | 0.136 | 0.34 | < 0.01 | 0.026 | 0.48 | [6 empty] | 11
| 1 | 6 | < 0.01 | 0.02 | < 1 | 0.1 | < 1 | < 0.1 | < 5 | 0.513 | 0.01 | 0.03 | < 0.01 | 0.06 | < 1 | 0.11 | 0.01 | 0.07 | 0.34 | 0.04 | < 0.08 | 0.04 | 0.132 | 0.33 | < 0.01 | 0.024 | 0.46
| Mean |  |  | 0.021667 |  | 0.068333 |  |  |  | 0.54883 | 0.025 | 0.026667 |  | 0.036667 |  | 0.10667 | 0.0125 | 0.073333 | 0.1825 | 0.038333 |  | 0.04 | 0.13617 | 0.33333 |  | 0.024667 | 0.495
| Median |  |  | 0.02 |  | 0.075 |  |  |  | 0.553 | 0.025 | 0.025 |  | 0.03 |  | 0.11 | 0.0125 | 0.07 | 0.145 | 0.04 |  | 0.04 | 0.136 | 0.33 |  | 0.024 | 0.49
| Std Dev. |  |  | 0.0040825 |  | 0.033116 |  |  |  | 0.019823 | 0.013784 | 0.008165 |  | 0.015055 |  | 0.013663 | 0.0020736 | 0.010328 | 0.1072 | 0.0040825 |  | 0 | 0.0076004 | 0.005164 |  | 0.0015055 | 0.037815
| Rel.Std.Dev. |  |  | 18.84% |  | 48.46% |  |  |  | 3.612% | 55.14% | 30.62% |  | 41.06% |  | 12.81% | 16.59% | 14.08% | 58.74% | 10.65% |  | 0% | 5.582% | 1.549% |  | 6.104% | 7.639%
| PDM3 |  |  | -61.94% |  | 20.05% |  |  |  | 864.2% | -56.08% | -53.15% |  | -35.59% |  | 87.39% | -78.04% | 28.83% | 220.6% | -32.66% |  | -29.73% | 139.2% | 485.6% |  | -56.67% | 769.6%
| Z-Score (Absolute) |  | 0.67 | 0.42 | 6.74 | 0.27 | 6.74 | 0 | 36.71 | 7.47 | 0.37 | 0.35 | 0.67 | 0.2 | 6.74 | 0.85 | 0.56 | 0.35 | 1.32 | 0.17 | 0.15 | 0.15 | 1.29 | 4.25 | 0.67 | 0.38 | 6.67
| Analytical results for Ho in OREAS 241b (Certified Value 0.42 ppm) | [32 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 0.452 | 0.402 | 0.3582 | 0.5 | 0.53 | 0.35 | [25 empty] | 1
| 1 | 2 | 0.445 | 0.427 | 0.37342 | 0.5 | 0.53 | 0.35 | [25 empty] | 19
| 1 | 3 | 0.445 | 0.389 | 0.35062 | 0.5 | 0.54 | 0.35 | [25 empty] | 16
| 1 | 4 | 0.452 | 0.417 | 0.39749 | 0.5 | 0.55 | 0.35 | [25 empty] | 0.4233
| 1 | 5 | 0.449 | 0.414 | 0.37264 | 0.5 | 0.54 | 0.35 | [25 empty] | 95
| 1 | 6 | 0.436 | 0.42 | 0.36858 | 0.5 | 0.54 | 0.35
| Mean |  | 0.4465 | 0.4115 | 0.37016 | 0.5 | 0.53833 | 0.35
| Median |  | 0.447 | 0.4155 | 0.37061 | 0.5 | 0.54 | 0.35
| Std Dev. |  | 0.0060249 | 0.013751 | 0.01607 | 0 | 0.0075277 | 6.0809e-17
| Rel.Std.Dev. |  | 1.349% | 3.342% | 4.341% | 0% | 1.398% | 1.737e-14%
| PDM3 |  | 5.481% | -2.787% | -12.55% | 18.12% | 27.18% | -17.32%
| Z-Score (Absolute) |  | 0.57 | 0 | 0.67 | NA | 2.07 | 1
| Analytical results for In in OREAS 241b (Certified Value 0.039 ppm) | [32 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 3
| 1 | 1 | 0.04 | 0.04 | 0.031104 | 0.04 | 0.05 | 0.04 | < 5 | 0.04 | 0.039 | 0.038 | 0.038 | 0.037 | < 10 | < 0.05 | 0.59 | < 0.1 | 0.04216 | 0.038 | 0.034 | 0.04 | 0.039 | 0.032 | [9 empty] | 1
| 1 | 2 | 0.03 | 0.04 | 0.031137 | 0.04 | 0.05 | 0.04 | < 5 | 0.04 | 0.038 | 0.039 | 0.033 | 0.038 | < 10 | < 0.05 | 0.56 | < 0.1 | 0.04032 | 0.041 | 0.042 | 0.03 | 0.037 | 0.031 | [9 empty] | 20
| 1 | 3 | 0.03 | 0.04 | 0.032561 | 0.03 | < 0.05 | 0.04 | < 5 | 0.05 | 0.037 | 0.037 | 0.036 | 0.037 | < 10 | < 0.05 | 0.76 | < 0.1 | 0.03981 | 0.041 | 0.042 | 0.04 | 0.038 | 0.034 | [9 empty] | 16
| 1 | 4 | 0.04 | 0.04 | 0.033654 | 0.04 | 0.05 | 0.04 | < 5 | 0.04 | 0.035 | 0.04 | 0.039 | 0.039 | < 10 | < 0.05 | 1 | < 0.1 | 0.04034 | 0.039 | 0.042 | 0.04 | 0.039 | 0.033 | [9 empty] | 0.038595
| 1 | 5 | 0.04 | 0.04 | 0.030574 | 0.05 | 0.05 | 0.04 | < 5 | 0.04 | 0.04 | 0.041 | 0.035 | 0.038 | < 10 | < 0.05 | < 0.1 | < 0.1 | 0.04004 | 0.04 | 0.036 | 0.04 | 0.039 | 0.031 | [9 empty] | 96
| 1 | 6 | 0.04 | 0.04 | 0.032992 | 0.04 | 0.05 | 0.04 | < 5 | 0.04 | 0.038 | 0.04 | 0.034 | 0.037 | < 10 | < 0.05 | 0.72 | < 0.1 | 0.04111 | 0.039 | 0.036 | 0.04 | 0.037 | 0.031
| Mean |  | 0.036667 | 0.04 | 0.032004 | 0.04 | 0.05 | 0.04 |  | 0.041667 | 0.037833 | 0.039167 | 0.035833 | 0.037667 |  |  | 0.726 |  | 0.04063 | 0.039667 | 0.038667 | 0.038333 | 0.038167 | 0.032
| Median |  | 0.04 | 0.04 | 0.031849 | 0.04 | 0.05 | 0.04 |  | 0.04 | 0.038 | 0.0395 | 0.0355 | 0.0375 |  |  | 0.72 |  | 0.04033 | 0.0395 | 0.039 | 0.04 | 0.0385 | 0.0315
| Std Dev. |  | 0.005164 | 0 | 0.0012343 | 0.0063246 | 0 | 0 |  | 0.0040825 | 0.0017224 | 0.001472 | 0.0023166 | 0.0008165 |  |  | 0.17487 |  | 0.00086856 | 0.0012111 | 0.0037238 | 0.0040825 | 0.00098319 | 0.0012649
| Rel.Std.Dev. |  | 14.08% | 0% | 3.857% | 15.81% | 0% | 0% |  | 9.798% | 4.553% | 3.758% | 6.465% | 2.168% |  |  | 24.09% |  | 2.138% | 3.053% | 9.631% | 10.65% | 2.576% | 3.953%
| PDM3 |  | -4.995% | 3.642% | -17.08% | 3.642% | 29.55% | 3.642% |  | 7.96% | -1.972% | 1.482% | -7.154% | -2.404% |  |  | 1781% |  | 5.274% | 2.778% | 0.1868% | -0.6769% | -1.109% | -17.09%
| Z-Score (Absolute) |  | 0.93 | 0.2 | 2.5 | 0.2 | 2.16 | 0.2 | 829.6 | 0.76 | 0.53 | 0.08 | 1.21 | 0.59 | 1672.48 | 4.86 | 193.5 | 3.57 | 0.41 | 0.08 | 0.25 | 0.37 | 0.42 | 2.5
| Analytical results for Ir in OREAS 241b (Indicative Value < 0.003 ppm) | [32 empty] | Indicative
Ir | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.27 | 1.28 | [29 empty] | ppm
|  |  | AR*MS | AR*MS | [29 empty] | 3
|  |  | 25g | 0.5g | [29 empty] | 3
| 1 | 1 | N.A. | < 0.003 | [29 empty] | 1
| 1 | 2 | N.A. | < 0.003 | [29 empty] | 6
| 1 | 3 | N.A. | < 0.003 | [29 empty] | 16
| 1 | 4 | N.A. | < 0.003 | [29 empty] | < 0.003
| 1 | 5 | N.A. | < 0.003 | [29 empty] | 12
| 1 | 6 | N.A. | < 0.003
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 241b (Certified Value 0.183 wt.%) | [32 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.19 | 0.172 | 0.21598 | 0.18382 | 0.1719 | 0.15 | 0.1876 | 0.209 | 0.1817 | 0.18 | 0.19 | 0.18 | 0.18 | 0.17 | 0.19825 | 0.18 | 0.16 | 0.18612 | 0.191 | 0.18 | 0.19 | 0.18753 | 0.18405 | 0.1705 | 0.17 | 0.17 | 0.21 | 0.19 | 0.19 |  |  | 1
| 1 | 2 | 0.19 | 0.172 | 0.21796 | 0.18842 | 0.17904 | 0.16 | 0.1926 | 0.204 | 0.1867 | 0.18 | 0.2 | 0.18 | 0.18 | 0.17 | 0.2026 | 0.18 | 0.16 | 0.1856 | 0.189 | 0.18 | 0.18 | 0.19067 | 0.18235 | 0.1735 | 0.17 | 0.18 | 0.21 | 0.19 | 0.18 |  |  | #N/A
| 1 | 3 | 0.19 | 0.168 | 0.21678 | 0.18239 | 0.17399 | 0.17 | 0.1918 | 0.203 | 0.1823 | 0.19 | 0.19 | 0.18 | 0.18 | 0.17 | 0.19361 | 0.18 | 0.16 | 0.18932 | 0.188 | 0.19 | 0.18 | 0.18413 | 0.17994 | 0.176 | 0.18 | 0.18 | 0.2 | 0.19 | 0.19 |  |  | 16
| 1 | 4 | 0.19 | 0.171 | 0.21426 | 0.18401 | 0.17543 | 0.15 | 0.1893 | 0.202 | 0.1785 | 0.19 | 0.19 | 0.17 | 0.18 | 0.17 | 0.19846 | 0.18 | 0.16 | 0.18599 | 0.189 | 0.17 | 0.18 | 0.18671 | 0.16903 | 0.1685 | 0.17 | 0.18 | 0.2 | 0.19 | 0.18 |  |  | 0.18316
| 1 | 5 | 0.19 | 0.171 | 0.21649 | 0.18705 | 0.17693 | 0.15 | 0.1909 | 0.208 | 0.1729 | 0.18 | 0.19 | 0.18 | 0.18 | 0.18 | 0.18969 | 0.18 | 0.16 | 0.17926 | 0.191 | 0.18 | 0.19 | 0.18234 | 0.16954 | 0.175 | 0.17 | 0.18 | 0.21 | 0.19 | 0.19 |  |  | 97
| 1 | 6 | 0.19 | 0.169 | 0.21723 | 0.18966 | 0.17235 | 0.15 | 0.1876 | 0.215 | 0.1763 | 0.18 | 0.19 | 0.18 | 0.18 | 0.17 | 0.19641 | 0.18 | 0.16 | 0.17907 | 0.188 | 0.18 | 0.19 | 0.18236 | 0.18447 | 0.178 | 0.17 | 0.17 | 0.2 | 0.19 | 0.19
| Mean |  | 0.19 | 0.1705 | 0.21645 | 0.18589 | 0.17494 | 0.155 | 0.18997 | 0.20683 | 0.17973 | 0.18333 | 0.19167 | 0.17833 | 0.18 | 0.17167 | 0.1965 | 0.18 | 0.16 | 0.18423 | 0.18933 | 0.18 | 0.185 | 0.18562 | 0.17823 | 0.17358 | 0.17167 | 0.17667 | 0.205 | 0.19 | 0.18667
| Median |  | 0.19 | 0.171 | 0.21663 | 0.18553 | 0.17471 | 0.15 | 0.1901 | 0.206 | 0.1801 | 0.18 | 0.19 | 0.18 | 0.18 | 0.17 | 0.19733 | 0.18 | 0.16 | 0.18579 | 0.189 | 0.18 | 0.185 | 0.18542 | 0.18115 | 0.17425 | 0.17 | 0.18 | 0.205 | 0.19 | 0.19
| Std Dev. |  | 3.0405e-17 | 0.0016432 | 0.0012675 | 0.0028993 | 0.002752 | 0.0083666 | 0.002136 | 0.0048751 | 0.0048784 | 0.005164 | 0.0040825 | 0.0040825 | 3.0405e-17 | 0.0040825 | 0.0044449 | 3.0405e-17 | 0 | 0.0041435 | 0.0013663 | 0.0063246 | 0.0054772 | 0.0032856 | 0.0071115 | 0.0035414 | 0.0040825 | 0.005164 | 0.0054772 | 3.0405e-17 | 0.005164
| Rel.Std.Dev. |  | 1.6e-14% | 0.9637% | 0.5856% | 1.56% | 1.573% | 5.398% | 1.124% | 2.357% | 2.714% | 2.817% | 2.13% | 2.289% | 1.689e-14% | 2.378% | 2.262% | 1.689e-14% | 0% | 2.249% | 0.7216% | 3.514% | 2.961% | 1.77% | 3.99% | 2.04% | 2.378% | 2.923% | 2.672% | 1.6e-14% | 2.766%
| PDM3 |  | 3.733% | -6.913% | 18.17% | 1.49% | -4.488% | -15.38% | 3.715% | 12.92% | -1.872% | 0.09369% | 4.643% | -2.636% | -1.726% | -6.276% | 7.285% | -1.726% | -12.65% | 0.5819% | 3.369% | -1.726% | 1.004% | 1.343% | -2.693% | -5.229% | -6.276% | -3.546% | 11.92% | 3.733% | 1.914%
| Z-Score (Absolute) |  | 0.67 | 1.3 | 3.35 | 0.26 | 0.85 | 2.87 | 0.67 | 2.38 | 0.36 | 0 | 0.84 | 0.51 | 0.34 | 1.18 | 1.33 | 0.34 | 2.36 | 0.09 | 0.61 | 0.34 | 0.17 | 0.23 | 0.52 | 0.99 | 1.18 | 0.67 | 2.19 | 0.67 | 0.34
| Analytical results for La in OREAS 241b (Certified Value 5.11 ppm) | [32 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 4.9 | 5.136 | 5.136 | 5.53 | 5.4088 | 5.1 | 2 | 5.8 | 5.152 | < 5 | 5 | 4.9 | 5.1 | 4.8 | 5.1 | 5 | 5.41 | 4.5 | 5.3 | 5 | 5.5 | 5.3 | 5.3 | 5.4 | 5.2 | [6 empty] | 1
| 1 | 2 | 4.8 | 5.072 | 4.788 | 5.59 | 5.4351 | 5.3 | 2 | 5.67 | 5.269 | 5 | 5 | 4.9 | 5 | 4.8 | 5.1 | 5 | 5.43 | 4.7 | 5.1 | 5 | 5.5 | 5.5 | 4.9 | 5.2 | 5.2 | [6 empty] | 21
| 1 | 3 | 4.8 | 5.119 | 4.57 | 5.51 | 5.2398 | 5.3 | 2 | 5.73 | 5.217 | 5 | 5 | 4.9 | 5 | 4.8 | 5.1 | 5 | 5.35 | 5 | 5.3 | 5 | 5.4 | 5.4 | 5 | 5.1 | 5.3 | [6 empty] | 16
| 1 | 4 | 4.8 | 5.197 | 5.65 | 5.54 | 5.3462 | 5 | 2 | 5.71 | 5.093 | 5 | 5 | 4.8 | 5.2 | 4.9 | 5 | 5 | 5.06 | 4.3 | 5 | 5 | 5.2 | 5.5 | 5.1 | 5.2 | 5.1 | [6 empty] | 5.1149
| 1 | 5 | 4.8 | 5.138 | 4.921 | 5.65 | 5.5228 | 5.1 | 1 | 5.87 | 5.001 | 5 | 5 | 4.9 | 5.2 | 4.8 | 5 | 5 | 4.93 | 5 | 5 | 5 | 5.3 | 5.4 | 5.1 | 5.2 | 5.3 | [6 empty] | 98
| 1 | 6 | 4.8 | 5.059 | 4.772 | 5.59 | 5.3277 | 4.8 | 2 | 5.97 | 4.895 | 5 | 6 | 4.9 | 5.2 | 4.8 | 4.9 | 5 | 5.08 | 4.7 | 5 | 5 | 5.5 | 5.3 | 5 | 5.2 | 5.2
| Mean |  | 4.8167 | 5.1202 | 4.9728 | 5.5683 | 5.3801 | 5.1 | 1.8333 | 5.7917 | 5.1045 | 5 | 5.1667 | 4.8833 | 5.1167 | 4.8167 | 5.0333 | 5 | 5.21 | 4.7 | 5.1167 | 5 | 5.4 | 5.4 | 5.0667 | 5.2167 | 5.2167
| Median |  | 4.8 | 5.1275 | 4.8545 | 5.565 | 5.3775 | 5.1 | 2 | 5.765 | 5.1225 | 5 | 5 | 4.9 | 5.15 | 4.8 | 5.05 | 5 | 5.215 | 4.7 | 5.05 | 5 | 5.45 | 5.4 | 5.05 | 5.2 | 5.2
| Std Dev. |  | 0.040825 | 0.05007 | 0.38069 | 0.051543 | 0.097759 | 0.18974 | 0.40825 | 0.1125 | 0.13907 | 0 | 0.40825 | 0.040825 | 0.098319 | 0.040825 | 0.08165 | 0 | 0.21251 | 0.27568 | 0.1472 | 0 | 0.12649 | 0.089443 | 0.13663 | 0.098319 | 0.075277
| Rel.Std.Dev. |  | 0.8476% | 0.9779% | 7.655% | 0.9256% | 1.817% | 3.72% | 22.27% | 1.942% | 2.725% | 0% | 7.902% | 0.836% | 1.922% | 0.8476% | 1.622% | 0% | 4.079% | 5.866% | 2.877% | 0% | 2.342% | 1.656% | 2.697% | 1.885% | 1.443%
| PDM3 |  | -5.831% | 0.1022% | -2.778% | 8.864% | 5.183% | -0.292% | -64.16% | 13.23% | -0.204% | -2.247% | 1.011% | -4.528% | 0.03382% | -5.831% | -1.595% | -2.247% | 1.859% | -8.112% | 0.03382% | -2.247% | 5.573% | 5.573% | -0.9437% | 1.989% | 1.989%
| Z-Score (Absolute) |  | 1.66 | 0.02 | 0.8 | 2.5 | 1.46 | 0.09 | NA | 3.73 | 0.07 | NA | NA | 1.29 | 0 | 1.66 | 0.46 | NA | 0.52 | 2.3 | 0 | NA | 1.57 | 1.57 | 0.28 | 0.55 | 0.55
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution. Results from laboratories 1.08, 1.12, 1.13 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 241b (Certified Value 9.85 ppm) | [32 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | < 1 | 9.38 | 16.6 | 3.5655 | 9.4477 | 9.3 | 6 | 11 | 9.6 | 10 | 11.3 | 10.5 | 10 | 10.2 | 9.8 | < 5 | 8.7028 | 9 | 10 | 9 | 10.2 | 9.1 | 11 | 10.3 | 9.1 | [6 empty] | 1
| 1 | 2 | < 1 | 9.24 | 16.34 | 3.5453 | 9.7494 | 9.9 | 6 | 10.7 | 9.8 | 10 | 11 | 10.3 | 10 | 10.1 | 9.9 | < 5 | 8.7967 | 9.5 | 9 | 9 | 10.3 | 9.3 | 10 | 9.9 | 9.4 | [6 empty] | #N/A
| 1 | 3 | < 1 | 9.19 | 15.56 | 3.5616 | 9.48 | 9.3 | 6 | 11.2 | 9.8 | 10 | 10.8 | 10.3 | 9.9 | 10.2 | 9.8 | < 5 | 7.7355 | 9.6 | 10 | 9 | 10.4 | 9.4 | 10 | 9.9 | 9.2 | [6 empty] | 16
| 1 | 4 | < 1 | 9.23 | 15.96 | 3.6224 | 9.6167 | 9 | 6 | 10.7 | 9.5 | 10 | 10.8 | 10.2 | 10 | 10 | 9.6 | < 5 | 8.0616 | 9.3 | 9 | 9 | 10.1 | 9.3 | 10 | 10.5 | 9.2 | [6 empty] | 9.8497
| 1 | 5 | < 1 | 9.44 | 15.65 | 3.633 | 9.6063 | 9.2 | 6 | 11.1 | 9.3 | 10 | 11.1 | 10.2 | 10 | 10.4 | 9.6 | < 5 | 7.882 | 8.9 | 10 | 9 | 10.2 | 9.5 | 11 | 9.9 | 9.3 | [6 empty] | 99
| 1 | 6 | < 1 | 9.23 | 16.28 | 3.6037 | 9.3207 | 8.8 | 6 | 11.1 | 9.2 | 10 | 10.6 | 10.2 | 9.8 | 10.3 | 9.9 | < 5 | 7.1369 | 9.3 | 9 | 9 | 10.3 | 8.9 | 10 | 9.7 | 8.5
| Mean |  |  | 9.285 | 16.065 | 3.5886 | 9.5368 | 9.25 | 6 | 10.9667 | 9.5333 | 10 | 10.9333 | 10.2833 | 9.95 | 10.2 | 9.7667 |  | 8.0526 | 9.2667 | 9.5 | 9 | 10.25 | 9.25 | 10.3333 | 10.0333 | 9.1167
| Median |  |  | 9.235 | 16.12 | 3.5846 | 9.5432 | 9.25 | 6 | 11.05 | 9.55 | 10 | 10.9 | 10.25 | 10 | 10.2 | 9.8 |  | 7.9718 | 9.3 | 9.5 | 9 | 10.25 | 9.3 | 10 | 9.9 | 9.2
| Std Dev. |  |  | 0.10015 | 0.41142 | 0.035969 | 0.15118 | 0.37283 | 0 | 0.21602 | 0.25033 | 0 | 0.25033 | 0.1169 | 0.083666 | 0.14142 | 0.13663 |  | 0.62364 | 0.27325 | 0.54772 | 0 | 0.10488 | 0.21679 | 0.5164 | 0.30111 | 0.31885
| Rel.Std.Dev. |  |  | 1.079% | 2.561% | 1.002% | 1.585% | 4.031% | 0% | 1.97% | 2.626% | 0% | 2.29% | 1.137% | 0.8409% | 1.386% | 1.399% |  | 7.745% | 2.949% | 5.766% | 0% | 1.023% | 2.344% | 4.997% | 3.001% | 3.497%
| PDM3 |  |  | -5.733% | 63.1% | -63.57% | -3.177% | -6.088% | -39.08% | 11.34% | -3.212% | 1.526% | 11% | 4.403% | 1.019% | 3.557% | -0.8427% |  | -18.25% | -5.919% | -3.55% | -8.626% | 4.064% | -6.088% | 4.91% | 1.865% | -7.442%
| Z-Score (Absolute) |  | 10.48 | 0.29 | 7.57 | 6.89 | 0 | 0.33 | NA | 1.66 | 0 | NA | 1.62 | 0.87 | 0.48 | 0.77 | 0.27 | 8.16 | 1.72 | 0.31 | NA | NA | 0.83 | 0.33 | NA | 0.58 | 0.49
Results from laboratories 1.08, 1.12, 1.24, 1.27 and 1.32 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 241b (Certified Value 0.15 ppm) | [32 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | 1.32 | [24 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | [24 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [24 empty] | 2
| 1 | 1 | 0.14 | 0.162 | 0.137 | 0.83883 | 0.2 | 0.1 | 0.14 | [24 empty] | 1
| 1 | 2 | 0.14 | 0.161 | 0.141 | 0.83869 | 0.2 | 0.1 | 0.14 | [24 empty] | 22
| 1 | 3 | 0.13 | 0.158 | 0.131 | 0.84006 | 0.2 | 0.1 | 0.15 | [24 empty] | 16
| 1 | 4 | 0.13 | 0.164 | 0.141 | 0.85825 | 0.2 | 0.1 | 0.14 | [24 empty] | 0.14624
| 1 | 5 | 0.14 | 0.163 | 0.141 | 0.82268 | 0.2 | 0.1 | 0.14 | [24 empty] | 100
| 1 | 6 | 0.13 | 0.158 | 0.139 | 0.83818 | 0.2 | 0.1 | 0.14
| Mean |  | 0.135 | 0.161 | 0.13833 | 0.83945 | 0.2 | 0.1 | 0.14167
| Median |  | 0.135 | 0.1615 | 0.14 | 0.83876 | 0.2 | 0.1 | 0.14
| Std Dev. |  | 0.0054772 | 0.0025298 | 0.0039328 | 0.011291 | 3.0405e-17 | 1.5202e-17 | 0.0040825
| Rel.Std.Dev. |  | 4.057% | 1.571% | 2.843% | 1.345% | 1.52e-14% | 1.52e-14% | 2.882%
| PDM3 |  | -7.689% | 10.09% | -5.41% | 474% | 36.76% | -31.62% | -3.13%
| Z-Score (Absolute) |  | 0.23 | 0.67 | 0.12 | 24.34 | 2.03 | 1.45 | 0
| Analytical results for Mg in OREAS 241b (Certified Value 1.58 wt.%) | [32 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 1.63 | 1.548 | 1.5695 | 1.6467 | 1.5559 | 1.63 | 1.7308 | 1.75 | 1.58 | 1.56 | 1.51 | 1.58 | 1.53 | 1.48 | 1.6433 | 1.58 | 1.46 | 1.6321 | 1.63 | 1.51 | 1.58 | 2.097 | 1.4351 | 1.441 | 1.54 | 1.54 | 1.55 | 1.6 | 1.63 |  |  | 1
| 1 | 2 | 1.62 | 1.53 | 1.607 | 1.6767 | 1.6075 | 1.7 | 1.7429 | 1.72 | 1.62 | 1.55 | 1.5 | 1.53 | 1.52 | 1.5 | 1.6761 | 1.6 | 1.47 | 1.6651 | 1.65 | 1.55 | 1.56 | 2.1664 | 1.4691 | 1.469 | 1.54 | 1.58 | 1.55 | 1.57 | 1.6 |  |  | #N/A
| 1 | 3 | 1.62 | 1.521 | 1.5735 | 1.6619 | 1.5667 | 1.72 | 1.7248 | 1.74 | 1.56 | 1.58 | 1.5 | 1.55 | 1.53 | 1.49 | 1.578 | 1.59 | 1.46 | 1.6735 | 1.68 | 1.53 | 1.58 | 2.111 | 1.4484 | 1.4955 | 1.55 | 1.59 | 1.52 | 1.57 | 1.65 |  |  | 16
| 1 | 4 | 1.63 | 1.501 | 1.5703 | 1.6248 | 1.5906 | 1.66 | 1.7127 | 1.7 | 1.57 | 1.62 | 1.5 | 1.54 | 1.52 | 1.49 | 1.656 | 1.58 | 1.48 | 1.6949 | 1.62 | 1.45 | 1.57 | 1.9025 | 1.4338 | 1.4075 | 1.52 | 1.58 | 1.55 | 1.57 | 1.6 |  |  | 1.5763
| 1 | 5 | 1.66 | 1.52 | 1.5687 | 1.6536 | 1.598 | 1.65 | 1.7248 | 1.76 | 1.52 | 1.56 | 1.52 | 1.54 | 1.54 | 1.51 | 1.5581 | 1.56 | 1.46 | 1.6322 | 1.66 | 1.55 | 1.58 | 1.769 | 1.4374 | 1.4655 | 1.54 | 1.58 | 1.54 | 1.57 | 1.68 |  |  | 101
| 1 | 6 | 1.64 | 1.5 | 1.5979 | 1.6824 | 1.5538 | 1.53 | 1.7006 | 1.79 | 1.52 | 1.59 | 1.52 | 1.54 | 1.52 | 1.5 | 1.6506 | 1.6 | 1.46 | 1.6135 | 1.67 | 1.51 | 1.56 | 1.9337 | 1.5097 | 1.5035 | 1.54 | 1.55 | 1.55 | 1.55 | 1.66
| Mean |  | 1.6333 | 1.52 | 1.5811 | 1.6577 | 1.5788 | 1.6483 | 1.7228 | 1.7433 | 1.5617 | 1.5767 | 1.5083 | 1.5467 | 1.5267 | 1.495 | 1.627 | 1.585 | 1.465 | 1.6519 | 1.6517 | 1.5167 | 1.5717 | 1.9966 | 1.4556 | 1.4637 | 1.5383 | 1.57 | 1.5433 | 1.5717 | 1.6367
| Median |  | 1.63 | 1.5205 | 1.5719 | 1.6578 | 1.5787 | 1.655 | 1.7248 | 1.745 | 1.565 | 1.57 | 1.505 | 1.54 | 1.525 | 1.495 | 1.6469 | 1.585 | 1.46 | 1.6486 | 1.655 | 1.52 | 1.575 | 2.0154 | 1.4429 | 1.4672 | 1.54 | 1.58 | 1.55 | 1.57 | 1.64
| Std Dev. |  | 0.015055 | 0.018144 | 0.016828 | 0.021019 | 0.022922 | 0.066758 | 0.014635 | 0.031411 | 0.038166 | 0.02582 | 0.0098319 | 0.017512 | 0.008165 | 0.010488 | 0.047381 | 0.015166 | 0.0083666 | 0.030819 | 0.023166 | 0.037238 | 0.0098319 | 0.15273 | 0.029642 | 0.035497 | 0.0098319 | 0.02 | 0.012111 | 0.016021 | 0.03266
| Rel.Std.Dev. |  | 0.9218% | 1.194% | 1.064% | 1.268% | 1.452% | 4.05% | 0.8495% | 1.802% | 2.444% | 1.638% | 0.6518% | 1.132% | 0.5348% | 0.7015% | 2.912% | 0.9568% | 0.5711% | 1.866% | 1.403% | 2.455% | 0.6256% | 7.649% | 2.036% | 2.425% | 0.6391% | 1.274% | 0.7847% | 1.019% | 1.996%
| PDM3 |  | 3.615% | -3.574% | 0.3039% | 5.16% | 0.1531% | 4.567% | 9.289% | 10.59% | -0.931% | 0.02055% | -4.314% | -1.883% | -3.151% | -5.16% | 3.214% | 0.5492% | -7.063% | 4.791% | 4.778% | -3.786% | -0.2966% | 26.66% | -7.661% | -7.148% | -2.411% | -0.4024% | -2.094% | -0.2966% | 3.827%
| Z-Score (Absolute) |  | 0.75 | 0.63 | 0.12 | 1.05 | 0.09 | 0.93 | 1.84 | 2.09 | 0.12 | 0.06 | 0.77 | 0.3 | 0.55 | 0.93 | 0.67 | 0.16 | 1.3 | 0.98 | 0.97 | 0.67 | 0 | 5.18 | 1.41 | 1.32 | 0.41 | 0.02 | 0.35 | 0 | 0.79
| Analytical results for Mn in OREAS 241b (Certified Value 0.074 wt.%) | [32 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.0791 | 0.07658 | 0.063523 | 0.078576 | 0.07371 | 0.0767 | 0.0813 | 0.0794 | 0.0771 | 0.0686 | 0.0839 | 0.0715 | 0.0706 | 0.0694 | 0.074447 | 0.0733 | 0.0679 | 0.076839 | 0.0777 | 0.0705 | 0.0757 | 0.073823 | 0.079577 | 0.064 | 0.0691 | 0.0693 | 0.0785 | 0.0706 | 0.0767 |  |  | 1
| 1 | 2 | 0.0785 | 0.07572 | 0.065177 | 0.078719 | 0.075395 | 0.0802 | 0.0829 | 0.0776 | 0.0788 | 0.0686 | 0.0828 | 0.0692 | 0.0702 | 0.0702 | 0.075714 | 0.0738 | 0.0684 | 0.078566 | 0.0778 | 0.0728 | 0.0743 | 0.07437 | 0.078027 | 0.0645 | 0.0695 | 0.0689 | 0.0755 | 0.0692 | 0.076 |  |  | 23
| 1 | 3 | 0.0785 | 0.07522 | 0.064787 | 0.078104 | 0.073818 | 0.083 | 0.0821 | 0.0785 | 0.0757 | 0.0688 | 0.0842 | 0.0703 | 0.0695 | 0.0697 | 0.071518 | 0.0739 | 0.0681 | 0.076889 | 0.0787 | 0.0704 | 0.0752 | 0.072632 | 0.077855 | 0.067 | 0.0698 | 0.0699 | 0.0763 | 0.069 | 0.0771 |  |  | 16
| 1 | 4 | 0.0789 | 0.07527 | 0.063184 | 0.078674 | 0.074605 | 0.0764 | 0.0821 | 0.0772 | 0.077 | 0.069 | 0.0798 | 0.0692 | 0.0698 | 0.0698 | 0.074541 | 0.0735 | 0.0689 | 0.079551 | 0.0771 | 0.0678 | 0.0752 | 0.072912 | 0.077631 | 0.064 | 0.0688 | 0.069 | 0.0763 | 0.0695 | 0.0743 |  |  | 0.07382
| 1 | 5 | 0.0801 | 0.07559 | 0.064862 | 0.078673 | 0.075141 | 0.0816 | 0.0829 | 0.0796 | 0.073 | 0.0697 | 0.0838 | 0.0697 | 0.0698 | 0.0707 | 0.070512 | 0.0715 | 0.0678 | 0.073441 | 0.0788 | 0.0709 | 0.0754 | 0.072446 | 0.07787 | 0.065 | 0.0702 | 0.0703 | 0.078 | 0.0691 | 0.0772 |  |  | 102
| 1 | 6 | 0.0792 | 0.07431 | 0.063568 | 0.078822 | 0.072571 | 0.0748 | 0.0813 | 0.0789 | 0.0727 | 0.0691 | 0.0812 | 0.0689 | 0.0694 | 0.0703 | 0.074263 | 0.0737 | 0.0677 | 0.076056 | 0.078 | 0.0707 | 0.0745 | 0.071421 | 0.081841 | 0.067 | 0.0689 | 0.067 | 0.0745 | 0.0683 | 0.0773
| Mean |  | 0.07905 | 0.075448 | 0.064183 | 0.078595 | 0.074207 | 0.078783 | 0.0821 | 0.078533 | 0.075717 | 0.068967 | 0.082617 | 0.0698 | 0.069883 | 0.070017 | 0.073499 | 0.073283 | 0.068133 | 0.07689 | 0.078017 | 0.070517 | 0.07505 | 0.072934 | 0.0788 | 0.06525 | 0.069383 | 0.069067 | 0.076517 | 0.069283 | 0.076433
| Median |  | 0.079 | 0.07543 | 0.064177 | 0.078674 | 0.074212 | 0.07845 | 0.0821 | 0.0787 | 0.07635 | 0.0689 | 0.0833 | 0.06945 | 0.0698 | 0.07 | 0.074355 | 0.0736 | 0.068 | 0.076864 | 0.0779 | 0.0706 | 0.0752 | 0.072772 | 0.077948 | 0.06475 | 0.0693 | 0.06915 | 0.0763 | 0.06915 | 0.0769
| Std Dev. |  | 0.00059245 | 0.0007422 | 0.00085155 | 0.00025325 | 0.0010503 | 0.0032744 | 0.00071554 | 0.00096678 | 0.002431 | 0.00041312 | 0.0017623 | 0.00096747 | 0.00044907 | 0.00047081 | 0.0020157 | 0.00089981 | 0.00045019 | 0.0021204 | 0.00064317 | 0.0015993 | 0.00053944 | 0.0010461 | 0.0016476 | 0.0014053 | 0.00054924 | 0.0011466 | 0.0015052 | 0.00075741 | 0.0011483
| Rel.Std.Dev. |  | 0.7495% | 0.9837% | 1.327% | 0.3222% | 1.415% | 4.156% | 0.8715% | 1.231% | 3.211% | 0.599% | 2.133% | 1.386% | 0.6426% | 0.6724% | 2.742% | 1.228% | 0.6607% | 2.758% | 0.8244% | 2.268% | 0.7188% | 1.434% | 2.091% | 2.154% | 0.7916% | 1.66% | 1.967% | 1.093% | 1.502%
| PDM3 |  | 7.085% | 2.206% | -13.05% | 6.469% | 0.5245% | 6.724% | 11.22% | 6.386% | 2.57% | -6.574% | 11.92% | -5.445% | -5.332% | -5.152% | -0.4341% | -0.7264% | -7.703% | 4.16% | 5.686% | -4.474% | 1.667% | -1.2% | 6.747% | -11.61% | -6.01% | -6.439% | 3.654% | -6.145% | 3.541%
| Z-Score (Absolute) |  | 0.75 | 0.19 | 1.56 | 0.68 | 0 | 0.71 | 1.23 | 0.67 | 0.24 | 0.82 | 1.31 | 0.69 | 0.67 | 0.65 | 0.11 | 0.14 | 0.95 | 0.42 | 0.59 | 0.58 | 0.13 | 0.2 | 0.72 | 1.4 | 0.75 | 0.8 | 0.36 | 0.77 | 0.35
| Analytical results for Mo in OREAS 241b (Certified Value 2.16 ppm) | [32 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 3
| 1 | 1 | 2.49 | 2.16 | 2.1 | 1.5428 | 2.1086 | 2.12 | 1.97 | 2.2 | 2.1 | 2 | 2.1 | 2.21 | 2.14 | 2.08 | < 10 | 2.01 | < 2 | < 5 | 1.7 | 1.93 | 2.21 | 1.38 | 2 | 2.1816 | 2.21 | 2.094 | 2.33 | 2.34 | 2.19 | 2.14 |  | 1
| 1 | 2 | 2.14 | 2.16 | 2.29 | 1.5128 | 2.3068 | 2.18 | 1.91 | 2.2 | 2.2 | 2 | 2.1 | 2.22 | 2.22 | 2.02 | < 10 | 2.03 | < 2 | < 5 | 1.8 | 1.98 | 2.17 | 2.06 | 2 | 2.1977 | 2.18 | 2.239 | 2.29 | 2.22 | 2.26 | 2.11 |  | 24
| 1 | 3 | 2.19 | 2.17 | 1.91 | 1.528 | 2.2427 | 2.2 | 1.97 | 2.2 | 2.1 | 2 | 2.2 | 2.29 | 2.13 | 2.05 | < 10 | 1.96 | < 2 | < 5 | 1.8 | 1.97 | 2.14 | 2.11 | 2 | 2.2331 | 2.14 | 2.16 | 2.33 | 2.32 | 2.23 | 2.21 |  | 16
| 1 | 4 | 2.26 | 2.13 | 2.1 | 1.5011 | 2.2919 | 2.12 | 1.94 | 2.2 | 2 | 2 | 2.2 | 2.1 | 2.1 | 2.13 | < 10 | 1.96 | < 2 | < 5 | 1.9 | 1.83 | 2.21 | 2.19 | 2 | 2.1404 | 2.14 | 2.256 | 2.38 | 2.25 | 2.2 | 2.1 |  | 2.1628
| 1 | 5 | 2.2 | 2.18 | 2.11 | 1.6429 | 2.205 | 2.07 | 1.91 | 2.2 | 2.4 | 2 | 2.1 | 2.09 | 2.18 | 2.04 | < 10 | 1.99 | < 2 | < 5 | 1.8 | 1.96 | 2.26 | 2.18 | 2 | 2.1159 | 2.18 | 2.164 | 2.28 | 2.31 | 2.16 | 2.16 |  | 103
| 1 | 6 | 2.2 | 2.13 | 2.14 | 1.5904 | 2.2514 | 2.02 | 1.97 | 2.3 | 2.7 | 2 | 2.2 | 2.16 | 2.17 | 2.01 | < 10 | 1.97 | < 2 | < 5 | 1.9 | 1.9 | 2.18 | 2 | 2 | 2.0723 | 2.1 | 2.071 | 2.21 | 2.24 | 2.14 | 2.11
| Mean |  | 2.2467 | 2.155 | 2.1083 | 1.553 | 2.2344 | 2.1183 | 1.945 | 2.2167 | 2.25 | 2 | 2.15 | 2.1783 | 2.1567 | 2.055 |  | 1.9867 |  |  | 1.8167 | 1.9283 | 2.195 | 1.9867 | 2 | 2.1568 | 2.1583 | 2.164 | 2.3033 | 2.28 | 2.1967 | 2.1383
| Median |  | 2.2 | 2.16 | 2.105 | 1.5354 | 2.247 | 2.12 | 1.955 | 2.2 | 2.15 | 2 | 2.15 | 2.185 | 2.155 | 2.045 |  | 1.98 |  |  | 1.8 | 1.945 | 2.195 | 2.085 | 2 | 2.161 | 2.16 | 2.162 | 2.31 | 2.28 | 2.195 | 2.125
| Std Dev. |  | 0.12517 | 0.020736 | 0.12123 | 0.053897 | 0.071537 | 0.067057 | 0.029496 | 0.040825 | 0.25884 | 0 | 0.054772 | 0.07679 | 0.042269 | 0.044159 |  | 0.028752 |  |  | 0.075277 | 0.056362 | 0.041352 | 0.30579 | 0 | 0.058637 | 0.0392 | 0.074369 | 0.05785 | 0.049396 | 0.044121 | 0.041673
| Rel.Std.Dev. |  | 5.571% | 0.9622% | 5.75% | 3.471% | 3.202% | 3.166% | 1.516% | 1.842% | 11.5% | 0% | 2.548% | 3.525% | 1.96% | 2.149% |  | 1.447% |  |  | 4.144% | 2.923% | 1.884% | 15.39% | 0% | 2.719% | 1.816% | 3.437% | 2.512% | 2.167% | 2.009% | 1.949%
| PDM3 |  | 3.88% | -0.3587% | -2.516% | -28.19% | 3.313% | -2.054% | -10.07% | 2.493% | 4.034% | -7.525% | -0.5899% | 0.7202% | -0.2816% | -4.982% |  | -8.142% |  |  | -16% | -10.84% | 1.491% | -8.142% | -7.525% | -0.2744% | -0.2046% | 0.05746% | 6.5% | 5.421% | 1.568% | -1.129%
| Z-Score (Absolute) |  | 0.72 | 0.01 | 0.39 | 4.86 | 0.63 | 0.31 | 1.7 | 0.48 | 0.75 | NA | 0.05 | 0.17 | 0 | 0.82 | 22.88 | 1.37 | 9.31 | 2.76 | 2.74 | 1.84 | 0.31 | 1.37 | NA | 0 | 0.01 | 0.06 | 1.18 | 0.99 | 0.32 | 0.15
Results from laboratories 1.12 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Na in OREAS 241b (Certified Value 0.283 wt.%) | [32 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.3 | 0.288 | 1.275 | 0.27844 | 0.27091 | 0.265 | 0.289 | 0.27 | 0.27 | 0.26 | 0.28 | 0.3 | 0.29 | 0.29 | 0.36614 | 0.29 | 0.26 | 0.28186 | 0.3 | 0.29 | 0.29 | 0.34 | 0.30677 | 0.2725 | 0.28 | 0.27 | 0.3 | 0.31 | 0.35 |  |  | 1
| 1 | 2 | 0.3 | 0.288 | 1.2569 | 0.27491 | 0.28025 | 0.27 | 0.297 | 0.26 | 0.28 | 0.26 | 0.28 | 0.29 | 0.29 | 0.29 | 0.36584 | 0.28 | 0.27 | 0.28299 | 0.29 | 0.304 | 0.28 | 0.33 | 0.31699 | 0.275 | 0.27 | 0.28 | 0.33 | 0.3 | 0.34 |  |  | #N/A
| 1 | 3 | 0.3 | 0.286 | 1.1271 | 0.27836 | 0.27567 | 0.278 | 0.297 | 0.27 | 0.26 | 0.26 | 0.28 | 0.29 | 0.29 | 0.29 | 0.35869 | 0.28 | 0.27 | 0.28432 | 0.29 | 0.309 | 0.29 | 0.32 | 0.3115 | 0.28 | 0.28 | 0.28 | 0.32 | 0.3 | 0.35 |  |  | 16
| 1 | 4 | 0.3 | 0.279 | 1.2482 | 0.2777 | 0.2751 | 0.264 | 0.297 | 0.26 | 0.27 | 0.27 | 0.28 | 0.29 | 0.29 | 0.29 | 0.3486 | 0.28 | 0.27 | 0.28685 | 0.29 | 0.283 | 0.28 | 0.32 | 0.29205 | 0.2665 | 0.27 | 0.28 | 0.33 | 0.3 | 0.33 |  |  | 0.28305
| 1 | 5 | 0.31 | 0.287 | 1.1113 | 0.27793 | 0.27952 | 0.276 | 0.297 | 0.27 | 0.26 | 0.27 | 0.29 | 0.29 | 0.3 | 0.29 | 0.34364 | 0.28 | 0.27 | 0.27032 | 0.3 | 0.311 | 0.28 | 0.31 | 0.29326 | 0.2795 | 0.28 | 0.27 | 0.35 | 0.3 | 0.34 |  |  | 104
| 1 | 6 | 0.31 | 0.283 | 1.3647 | 0.27042 | 0.27258 | 0.254 | 0.289 | 0.26 | 0.26 | 0.27 | 0.28 | 0.29 | 0.29 | 0.29 | 0.35414 | 0.29 | 0.27 | 0.27546 | 0.29 | 0.3 | 0.28 | 0.32 | 0.29653 | 0.2845 | 0.27 | 0.27 | 0.33 | 0.3 | 0.35
| Mean |  | 0.30333 | 0.28517 | 1.2305 | 0.27629 | 0.27567 | 0.26783 | 0.29433 | 0.265 | 0.26667 | 0.265 | 0.28167 | 0.29167 | 0.29167 | 0.29 | 0.35618 | 0.28333 | 0.26833 | 0.2803 | 0.29333 | 0.2995 | 0.28333 | 0.32333 | 0.30285 | 0.27633 | 0.275 | 0.275 | 0.32667 | 0.30167 | 0.34333
| Median |  | 0.3 | 0.2865 | 1.2525 | 0.27782 | 0.27539 | 0.2675 | 0.297 | 0.265 | 0.265 | 0.265 | 0.28 | 0.29 | 0.29 | 0.29 | 0.35642 | 0.28 | 0.27 | 0.28243 | 0.29 | 0.302 | 0.28 | 0.32 | 0.30165 | 0.27725 | 0.275 | 0.275 | 0.33 | 0.3 | 0.345
| Std Dev. |  | 0.005164 | 0.0035449 | 0.095837 | 0.0031606 | 0.0036983 | 0.0088185 | 0.0041312 | 0.0054772 | 0.008165 | 0.0054772 | 0.0040825 | 0.0040825 | 0.0040825 | 0 | 0.0091393 | 0.005164 | 0.0040825 | 0.006192 | 0.005164 | 0.011005 | 0.005164 | 0.010328 | 0.010378 | 0.006377 | 0.0054772 | 0.0054772 | 0.01633 | 0.0040825 | 0.008165
| Rel.Std.Dev. |  | 1.702% | 1.243% | 7.788% | 1.144% | 1.342% | 3.293% | 1.404% | 2.067% | 3.062% | 2.067% | 1.449% | 1.4% | 1.4% | 0% | 2.566% | 1.823% | 1.521% | 2.209% | 1.76% | 3.674% | 1.823% | 3.194% | 3.427% | 2.308% | 1.992% | 1.992% | 4.999% | 1.353% | 2.378%
| PDM3 |  | 7.166% | 0.7477% | 334.7% | -2.387% | -2.607% | -5.376% | 3.986% | -6.377% | -5.788% | -6.377% | -0.4889% | 3.044% | 3.044% | 2.455% | 25.83% | 0.09997% | -5.199% | -0.9719% | 3.633% | 5.812% | 0.09997% | 14.23% | 6.995% | -2.373% | -2.844% | -2.844% | 15.41% | 6.577% | 21.3%
| Z-Score (Absolute) |  | 1.2 | 0 | 62.7 | 0.59 | 0.63 | 1.15 | 0.6 | 1.34 | 1.23 | 1.34 | 0.23 | 0.43 | 0.43 | 0.32 | 4.71 | 0.12 | 1.12 | 0.32 | 0.54 | 0.95 | 0.12 | 2.53 | 1.17 | 0.59 | 0.67 | 0.67 | 2.75 | 1.09 | 3.86
| Analytical results for Nb in OREAS 241b (Indicative Value 0.12 ppm) | [32 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [9 empty] | 2
| 1 | 1 | < 0.05 | 0.16 | 33.72 | 0.17919 | < 0.1 | 0.36 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.12 | 0.16 | 0.1 | 0.09 | < 10 | < 0.2 | < 1 | 0.12 | 0.05 | 0.25 | 0.08 | 0.16 | [9 empty] | 1
| 1 | 2 | < 0.05 | 0.17 | 33.04 | 0.17705 | 0.1 | 0.35 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.11 | 0.17 | 0.1 | 0.08 | < 10 | < 0.2 | < 1 | 0.13 | < 0.05 | 0.23 | 0.07 | 0.15 | [9 empty] | 7
| 1 | 3 | < 0.05 | 0.16 | 36.51 | 0.49032 | 0.1 | 0.39 | < 0.2 | 0.08 | < 5 | < 0.1 | 0.13 | 0.16 | 0.1 | 0.09 | < 10 | < 0.2 | < 1 | 0.14 | < 0.05 | 0.22 | 0.07 | 0.15 | [9 empty] | 16
| 1 | 4 | < 0.05 | 0.17 | 32.68 | 0.15189 | < 0.1 | 0.36 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.12 | 0.16 | 0.11 | 0.08 | < 10 | < 0.2 | < 1 | 0.12 | < 0.05 | 0.24 | 0.07 | 0.14 | [9 empty] | 0.12471
| 1 | 5 | < 0.05 | 0.15 | 35.76 | 0.16412 | 0.1 | 0.39 | < 0.2 | 0.09 | < 5 | < 0.1 | 0.12 | 0.18 | 0.1 | 0.07 | < 10 | < 0.2 | < 1 | 0.14 | < 0.05 | 0.23 | 0.08 | 0.15 | [9 empty] | 13
| 1 | 6 | < 0.05 | 0.16 | 33.36 | 0.15914 | 0.1 | 0.38 | < 0.2 | 0.1 | < 5 | < 0.1 | 0.12 | 0.16 | 0.09 | 0.08 | < 10 | < 0.2 | < 1 | 0.13 | < 0.05 | 0.21 | 0.08 | 0.15
| Mean |  |  | 0.16167 | 34.1783 | 0.22028 | 0.1 | 0.37167 |  | 0.091667 |  |  | 0.12 | 0.165 | 0.1 | 0.081667 |  |  |  | 0.13 | 0.05 | 0.23 | 0.075 | 0.15
| Median |  |  | 0.16 | 33.54 | 0.17058 | 0.1 | 0.37 |  | 0.09 |  |  | 0.12 | 0.16 | 0.1 | 0.08 |  |  |  | 0.13 | 0.05 | 0.23 | 0.075 | 0.15
| Std Dev. |  |  | 0.0075277 | 1.5722 | 0.1327 | 0 | 0.017224 |  | 0.0075277 |  |  | 0.0063246 | 0.0083666 | 0.0063246 | 0.0075277 |  |  |  | 0.0089443 |  | 0.014142 | 0.0054772 | 0.0063246
| Rel.Std.Dev. |  |  | 4.656% | 4.6% | 60.24% | 0% | 4.634% |  | 8.212% |  |  | 5.27% | 5.071% | 6.325% | 9.218% |  |  |  | 6.88% |  | 6.149% | 7.303% | 4.216%
| PDM3 |  |  | 29.63% | 27305% | 76.63% | -19.82% | 198% |  | -26.5% |  |  | -3.779% | 32.3% | -19.82% | -34.52% |  |  |  | 4.239% | -59.91% | 84.42% | -39.86% | 20.28%
| Z-Score (Absolute) |  | 1.44 | 0.53 | 492.05 | 1.38 | 0.6 | 3.56 | 0.36 | 0.48 | 34.32 | 1.08 | 0.07 | 0.58 | 0.36 | 0.63 | 70.44 | 0.36 | 5.42 | 0.07 | 1.38 | 1.52 | 0.72 | 0.36
| Analytical results for Nd in OREAS 241b (Certified Value 6.52 ppm) | [32 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 6.901 | 6.768 | 5.2911 | 6.41 | 6.53 | 6 | [25 empty] | 1
| 1 | 2 | 6.838 | 6.426 | 5.3129 | 7.01 | 6.64 | 6 | [25 empty] | 26
| 1 | 3 | 6.886 | 6.067 | 5.2552 | 6.6 | 6.54 | 6 | [25 empty] | 16
| 1 | 4 | 6.956 | 7.006 | 5.3151 | 6.5 | 6.51 | 6 | [25 empty] | 6.5174
| 1 | 5 | 6.968 | 6.47 | 5.3308 | 6.63 | 6.6 | 6 | [25 empty] | 105
| 1 | 6 | 6.782 | 6.445 | 5.3181 | 6.49 | 6.55 | 7
| Mean |  | 6.8885 | 6.5303 | 5.3039 | 6.6067 | 6.5617 | 6.1667
| Median |  | 6.8935 | 6.4575 | 5.314 | 6.55 | 6.545 | 6
| Std Dev. |  | 0.070614 | 0.32225 | 0.027082 | 0.21304 | 0.048751 | 0.40825
| Rel.Std.Dev. |  | 1.025% | 4.935% | 0.5106% | 3.225% | 0.743% | 6.62%
| PDM3 |  | 5.693% | 0.1979% | -18.62% | 1.369% | 0.6787% | -5.382%
| Z-Score (Absolute) |  | 1.15 | 0.05 | 4.16 | 0.2 | 0.05 | 1.27
| Analytical results for Ni in OREAS 241b (Certified Value 55 ppm) | [32 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 1
| 1 | 1 | 55.2 | 54.8 | 46.97 | 53.48 | 56.232 | 51.3 | 54.3 | 60 | 55.1 | 49 | 48.8 | 56 | 52.3 | 54 | 57.432 | 55 | 47 | 57.5772 | 56 | 56.4 | 56.8 | 57.57 | 53.3048 | 51 | 57.8156 | 54.1 | 51.7 | 59.8 | 52.1 | 55 |  | 1
| 1 | 2 | 55.5 | 54.4 | 47.75 | 53.74 | 56.7755 | 54.8 | 56.2 | 60 | 57.4 | 49 | 47.3 | 55.8 | 51.2 | 55.5 | 59.077 | 55.8 | 47 | 56.686 | 59 | 57.4 | 54.2 | 56.77 | 51.5936 | 52 | 59.6603 | 55.1 | 52.9 | 57.6 | 51.2 | 55.3 |  | 27
| 1 | 3 | 55.1 | 53.6 | 47.37 | 52.28 | 55.8562 | 54.9 | 56.1 | 63 | 55.2 | 49 | 48.9 | 56.1 | 51.8 | 54.4 | 54.648 | 55.1 | 47 | 57.4992 | 60 | 56.3 | 54.4 | 54.43 | 52.4492 | 54 | 61.5521 | 55.2 | 54 | 59.2 | 49.5 | 56.5 |  | 16
| 1 | 4 | 55.8 | 53.8 | 46.09 | 52.92 | 57.218 | 54.6 | 56.4 | 60 | 53.8 | 49 | 48 | 55.2 | 52.5 | 55.2 | 56.9 | 54.1 | 48 | 59.3597 | 53 | 54.6 | 54.6 | 55.8 | 50.4252 | 51 | 60.8738 | 53.7 | 52 | 58.9 | 50.3 | 54.5 |  | 54.7118
| 1 | 5 | 56.4 | 54 | 47.43 | 53.82 | 56.7267 | 54.8 | 54 | 62 | 54.6 | 49 | 48.2 | 56.4 | 52.9 | 54.9 | 54.303 | 54.1 | 47 | 56.146 | 53 | 56.4 | 54.4 | 54.99 | 50.0204 | 53 | 58.1951 | 54.9 | 54.1 | 60.4 | 49.7 | 56.1 |  | 106
| 1 | 6 | 56.1 | 53.2 | 45.74 | 53.54 | 55.7529 | 50.9 | 55.9 | 59 | 56.8 | 50 | 49.2 | 55.9 | 52.6 | 54 | 57.06 | 54.5 | 47 | 56.1856 | 56 | 55.2 | 52.7 | 54.54 | 56.158 | 54 | 56.2145 | 54.5 | 52.4 | 58 | 50 | 54.8
| Mean |  | 55.6833 | 53.9667 | 46.8917 | 53.2967 | 56.4269 | 53.55 | 55.4833 | 60.6667 | 55.4833 | 49.1667 | 48.4 | 55.9 | 52.2167 | 54.6667 | 56.57 | 54.7667 | 47.1667 | 57.2423 | 56.1667 | 56.05 | 54.5167 | 55.6833 | 52.3252 | 52.5 | 59.0519 | 54.5833 | 52.85 | 58.9833 | 50.4667 | 55.3667
| Median |  | 55.65 | 53.9 | 47.17 | 53.51 | 56.4794 | 54.7 | 56 | 60 | 55.15 | 49 | 48.5 | 55.95 | 52.4 | 54.65 | 56.98 | 54.75 | 47 | 57.0926 | 56 | 56.35 | 54.4 | 55.395 | 52.0214 | 52.5 | 58.9277 | 54.7 | 52.65 | 59.05 | 50.15 | 55.15
| Std Dev. |  | 0.51153 | 0.57155 | 0.8038 | 0.5897 | 0.57536 | 1.9045 | 1.0496 | 1.5055 | 1.36 | 0.40825 | 0.70143 | 0.4 | 0.61779 | 0.6314 | 1.8006 | 0.66232 | 0.40825 | 1.2074 | 2.9269 | 0.99549 | 1.3152 | 1.2738 | 2.2415 | 1.3784 | 2.0127 | 0.5947 | 1.0134 | 1.0591 | 0.99733 | 0.77889
| Rel.Std.Dev. |  | 0.9186% | 1.059% | 1.714% | 1.106% | 1.02% | 3.556% | 1.892% | 2.482% | 2.451% | 0.8303% | 1.449% | 0.7156% | 1.183% | 1.155% | 3.183% | 1.209% | 0.8655% | 2.109% | 5.211% | 1.776% | 2.412% | 2.288% | 4.284% | 2.626% | 3.408% | 1.09% | 1.918% | 1.796% | 1.976% | 1.407%
| PDM3 |  | 1.776% | -1.362% | -14.29% | -2.586% | 3.135% | -2.123% | 1.41% | 10.88% | 1.41% | -10.14% | -11.54% | 2.172% | -4.56% | -0.08246% | 3.396% | 0.1003% | -13.79% | 4.625% | 2.659% | 2.446% | -0.3566% | 1.776% | -4.362% | -4.043% | 7.933% | -0.2348% | -3.403% | 7.807% | -7.759% | 1.197%
| Z-Score (Absolute) |  | 0.41 | 0.32 | 3.34 | 0.61 | 0.73 | 0.5 | 0.33 | 2.54 | 0.33 | 2.37 | 2.7 | 0.5 | 1.07 | 0.02 | 0.79 | 0.02 | 3.22 | 1.08 | 0.62 | 0.57 | 0.09 | 0.41 | 1.02 | 0.95 | 1.85 | 0.06 | 0.8 | 1.82 | 1.81 | 0.28
| Analytical results for P in OREAS 241b (Certified Value 0.043 wt.%) | [32 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.0446 | 0.0443 | 0.035078 | 0.042369 | 0.04 | 0.0415 | 0.0448 | 0.0422 | 0.0424 | 0.046 | 0.044 | 0.044 | 0.041 | 0.047341 | 0.043 | 0.0392 | 0.045448 | 0.045 | 0.041 | 0.04 | 0.0415 | 0.042 | 0.042 | 0.0451 | 0.0429 | 0.044 | [5 empty] | 1
| 1 | 2 | 0.0448 | 0.0429 | 0.036607 | 0.043261 | 0.041 | 0.0415 | 0.0437 | 0.041 | 0.0422 | 0.045 | 0.043 | 0.044 | 0.04 | 0.04854 | 0.043 | 0.0396 | 0.045372 | 0.045 | 0.043 | 0.04 | 0.0425 | 0.042 | 0.044 | 0.045 | 0.0421 | 0.044 | [5 empty] | #N/A
| 1 | 3 | 0.0453 | 0.0424 | 0.036366 | 0.042218 | 0.041 | 0.0419 | 0.0445 | 0.0395 | 0.0423 | 0.045 | 0.043 | 0.043 | 0.041 | 0.04723 | 0.043 | 0.039 | 0.045282 | 0.044 | 0.041 | 0.04 | 0.0435 | 0.043 | 0.043 | 0.0451 | 0.0423 | 0.044 | [5 empty] | 16
| 1 | 4 | 0.0452 | 0.043 | 0.035352 | 0.043457 | 0.041 | 0.0419 | 0.0437 | 0.0406 | 0.0434 | 0.045 | 0.043 | 0.043 | 0.041 | 0.05148 | 0.043 | 0.0397 | 0.046999 | 0.044 | 0.04 | 0.04 | 0.0425 | 0.042 | 0.043 | 0.0456 | 0.0421 | 0.043 | [5 empty] | 0.042972
| 1 | 5 | 0.0462 | 0.044 | 0.036628 | 0.043286 | 0.04 | 0.0415 | 0.0449 | 0.0391 | 0.0429 | 0.042 | 0.043 | 0.044 | 0.041 | 0.045852 | 0.042 | 0.0391 | 0.043977 | 0.045 | 0.041 | 0.04 | 0.043 | 0.043 | 0.044 | 0.0446 | 0.042 | 0.044 | [5 empty] | 107
| 1 | 6 | 0.0458 | 0.0445 | 0.035649 | 0.042517 | 0.04 | 0.0415 | 0.0453 | 0.0401 | 0.0427 | 0.046 | 0.043 | 0.043 | 0.041 | 0.049139 | 0.043 | 0.0396 | 0.043922 | 0.044 | 0.041 | 0.04 | 0.0435 | 0.042 | 0.044 | 0.0444 | 0.0417 | 0.044
| Mean |  | 0.045317 | 0.043517 | 0.035947 | 0.042851 | 0.0405 | 0.041633 | 0.044483 | 0.040417 | 0.04265 | 0.044833 | 0.043167 | 0.0435 | 0.040833 | 0.048264 | 0.042833 | 0.039367 | 0.045167 | 0.0445 | 0.041167 | 0.04 | 0.04275 | 0.042333 | 0.043333 | 0.044967 | 0.042183 | 0.043833
| Median |  | 0.04525 | 0.0435 | 0.036007 | 0.042889 | 0.0405 | 0.0415 | 0.04465 | 0.04035 | 0.04255 | 0.045 | 0.043 | 0.0435 | 0.041 | 0.047941 | 0.043 | 0.0394 | 0.045327 | 0.0445 | 0.041 | 0.04 | 0.04275 | 0.042 | 0.0435 | 0.04505 | 0.0421 | 0.044
| Std Dev. |  | 0.00060139 | 0.0008612 | 0.00067423 | 0.00054194 | 0.00054772 | 0.00020656 | 0.00065853 | 0.0011161 | 0.00045056 | 0.001472 | 0.00040825 | 0.00054772 | 0.00040825 | 0.0019453 | 0.00040825 | 0.00030111 | 0.0011363 | 0.00054772 | 0.00098319 | 0 | 0.00075829 | 0.0005164 | 0.0008165 | 0.00042269 | 0.00040208 | 0.00040825
| Rel.Std.Dev. |  | 1.327% | 1.979% | 1.876% | 1.265% | 1.352% | 0.4961% | 1.48% | 2.761% | 1.056% | 3.283% | 0.9457% | 1.259% | 0.9998% | 4.031% | 0.9531% | 0.7649% | 2.516% | 1.231% | 2.388% | 0% | 1.774% | 1.22% | 1.884% | 0.94% | 0.9532% | 0.9314%
| PDM3 |  | 5.457% | 1.268% | -16.35% | -0.2801% | -5.752% | -3.114% | 3.518% | -5.946% | -0.7483% | 4.333% | 0.454% | 1.23% | -4.976% | 12.32% | -0.3217% | -8.389% | 5.108% | 3.557% | -4.2% | -6.915% | -0.5156% | -1.485% | 0.8419% | 4.643% | -1.834% | 2.005%
| Z-Score (Absolute) |  | 1.01 | 0.28 | 2.82 | 0 | 0.96 | 0.51 | 0.67 | 0.99 | 0.08 | 0.81 | 0.13 | 0.27 | 0.82 | 2.22 | 0 | 1.42 | 0.95 | 0.68 | 0.69 | 1.16 | 0.04 | 0.21 | 0.2 | 0.87 | 0.27 | 0.41
| Analytical results for Pb in OREAS 241b (Certified Value 37.4 ppm) | [32 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  | 2
| 1 | 1 | 40.3 | 37.4 | 34.81 | 35.35 | 36.0884 | 27.2 | 37 | 35.2 | 38.6 | 34 | 40.1 | 36.9 | 34.9 | 37.6 | 37.4 | 30 | 35.5295 | 32.9 | 34.01 | 38.5 | 41.63 | 38.9577 | 41 | 38.504 | 37.5 | 39.833 | 36.6 | 37.9 | 37.5 | 39.1 |  | 1
| 1 | 2 | 39.9 | 36.8 | 31.97 | 35.67 | 36.6629 | 28.6 | 37.9 | 34.1 | 38.7 | 34 | 39.8 | 36.9 | 34.2 | 37.4 | 38.2 | 31 | 35.2193 | 36.6 | 34.82 | 38.1 | 39.42 | 37.5906 | 40 | 37.6542 | 37.8 | 43.162 | 37 | 36.3 | 36.6 | 39.2 |  | 28
| 1 | 3 | 40.4 | 36.8 | 29.38 | 35.67 | 35.5253 | 28.6 | 36.9 | 35.2 | 38.2 | 35 | 40.9 | 37.1 | 34.6 | 38.4 | 37.3 | 31 | 34.5229 | 34.7 | 35.81 | 38.4 | 39.58 | 36.5304 | 41 | 38.2009 | 37.3 | 74.857 | 37 | 36.9 | 37.2 | 39.4 |  | 16
| 1 | 4 | 40.6 | 37.3 | 30.03 | 36.47 | 37.5277 | 27.9 | 37.6 | 34.5 | 38 | 35 | 40.4 | 36.4 | 35.5 | 38.1 | 37.1 | 32 | 35.75 | 35.2 | 32.34 | 38.2 | 40.82 | 35.2842 | 40 | 37.8928 | 36.9 | 43.545 | 37.1 | 36.4 | 37.7 | 38 |  | 37.4093
| 1 | 5 | 41.1 | 37.1 | 30 | 35.92 | 36.2446 | 28 | 37.1 | 35.1 | 36.7 | 35 | 40.3 | 36.6 | 35.4 | 36.7 | 36.3 | 32 | 35.6036 | 36.2 | 35.88 | 38.1 | 40.45 | 35.3028 | 42 | 38.3107 | 38 | 42.335 | 37.1 | 38 | 37 | 39.5 |  | 108
| 1 | 6 | 41.1 | 37.4 | 30.3 | 36.21 | 35.1575 | 28.2 | 37.1 | 36.2 | 36.5 | 35 | 41.4 | 37.5 | 35.7 | 37.1 | 37.5 | 31 | 35.7431 | 33.7 | 35.13 | 37.8 | 40.82 | 32.868 | 42 | 38.091 | 37.2 | 40.74 | 37.3 | 36.5 | 36.1 | 39
| Mean |  | 40.5667 | 37.1333 | 31.0817 | 35.8817 | 36.2011 | 28.0833 | 37.2667 | 35.05 | 37.7833 | 34.6667 | 40.4833 | 36.9 | 35.05 | 37.55 | 37.3 | 31.1667 | 35.3947 | 34.8833 | 34.665 | 38.1833 | 40.4533 | 36.089 | 41 | 38.1089 | 37.45 | 47.412 | 37.0167 | 37 | 37.0167 | 39.0333
| Median |  | 40.5 | 37.2 | 30.165 | 35.795 | 36.1665 | 28.1 | 37.1 | 35.15 | 38.1 | 35 | 40.35 | 36.9 | 35.15 | 37.5 | 37.35 | 31 | 35.5665 | 34.95 | 34.975 | 38.15 | 40.635 | 35.9166 | 41 | 38.146 | 37.4 | 42.7485 | 37.05 | 36.7 | 37.1 | 39.15
| Std Dev. |  | 0.47188 | 0.28048 | 2.0235 | 0.40686 | 0.84066 | 0.52313 | 0.39328 | 0.71764 | 0.95376 | 0.5164 | 0.57764 | 0.38471 | 0.58224 | 0.62849 | 0.61644 | 0.75277 | 0.46908 | 1.4247 | 1.3313 | 0.24833 | 0.83474 | 2.1128 | 0.89443 | 0.30298 | 0.40373 | 13.5202 | 0.23166 | 0.7642 | 0.59133 | 0.53914
| Rel.Std.Dev. |  | 1.163% | 0.7553% | 6.51% | 1.134% | 2.322% | 1.863% | 1.055% | 2.047% | 2.524% | 1.49% | 1.427% | 1.043% | 1.661% | 1.674% | 1.653% | 2.415% | 1.325% | 4.084% | 3.841% | 0.6504% | 2.063% | 5.854% | 2.182% | 0.795% | 1.078% | 28.52% | 0.6258% | 2.065% | 1.597% | 1.381%
| PDM3 |  | 8.44% | -0.7376% | -16.91% | -4.083% | -3.23% | -24.93% | -0.3812% | -6.307% | 1% | -7.331% | 8.217% | -1.361% | -6.307% | 0.3762% | -0.2921% | -16.69% | -5.385% | -6.752% | -7.336% | 2.069% | 8.137% | -3.529% | 9.599% | 1.87% | 0.1089% | 26.74% | -1.049% | -1.094% | -1.049% | 4.341%
| Z-Score (Absolute) |  | 1.72 | 0.06 | 2.87 | 0.55 | 0.39 | 4.32 | 0.12 | 0.95 | 0.37 | 1.14 | 1.68 | 0.06 | 0.95 | 0.26 | 0.14 | 2.83 | 0.78 | 1.03 | 1.14 | 0.56 | 1.66 | 0.45 | 1.93 | 0.53 | 0.21 | 5.03 | 0 | 0.01 | 0 | 0.98
| Analytical results for Pd in OREAS 241b (Indicative Value 10.6 ppb) | [32 empty] | Indicative
Pd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 1
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 1
| 1 | 1 | 11 | < 10 | 13 | < 50 | < 20 | 28 | [25 empty] | 1
| 1 | 2 | 15 | < 10 | 11 | < 50 | < 20 | 32 | [25 empty] | 1
| 1 | 3 | 11 | 15 | 13 | < 50 | 25.71 | 34 | [25 empty] | 16
| 1 | 4 | 12 | 12 | 10 | < 50 | < 20 | 31 | [25 empty] | 10.5917
| 1 | 5 | 11 | 13 | 11 | < 50 | < 20 | 29 | [25 empty] | 11
| 1 | 6 | 11 | 14 | < 10 | < 50 | < 20 | 36
| Mean |  | 11.8333 | 13.5 | 11.6 |  | 25.71 | 31.6667
| Median |  | 11 | 13.5 | 11 |  | 25.71 | 31.5
| Std Dev. |  | 1.6021 | 1.291 | 1.3416 |  |  | 3.0111
| Rel.Std.Dev. |  | 13.54% | 9.563% | 11.57% |  |  | 9.509%
| PDM3 |  | 11.72% | 27.46% | 9.52% |  | 142.7% | 199%
| Z-Score (Absolute) |  | 0.16 | 0.64 | 0.71 | 5.24 | 0.16 | 7.98
| Analytical results for Pr in OREAS 241b (Certified Value 1.42 ppm) | [32 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 3
| 1 | 1 | 1.467 | 1.391 | 1.1753 | 1.5 | 1.37 | 1.4 | [25 empty] | 1
| 1 | 2 | 1.455 | 1.355 | 1.1883 | 1.5 | 1.4 | 1.4 | [25 empty] | 29
| 1 | 3 | 1.471 | 1.233 | 1.1726 | 1.5 | 1.39 | 1.3 | [25 empty] | 16
| 1 | 4 | 1.475 | 1.517 | 1.1665 | 1.5 | 1.39 | 1.4 | [25 empty] | 1.4185
| 1 | 5 | 1.477 | 1.345 | 1.1762 | 1.5 | 1.36 | 1.4 | [25 empty] | 109
| 1 | 6 | 1.442 | 1.327 | 1.1691 | 1.4 | 1.39 | 1.5
| Mean |  | 1.4645 | 1.3613 | 1.1747 | 1.4833 | 1.3833 | 1.4
| Median |  | 1.469 | 1.35 | 1.1739 | 1.5 | 1.39 | 1.4
| Std Dev. |  | 0.013502 | 0.092809 | 0.007623 | 0.040825 | 0.015055 | 0.063246
| Rel.Std.Dev. |  | 0.9219% | 6.817% | 0.6489% | 2.752% | 1.088% | 4.518%
| PDM3 |  | 3.243% | -4.03% | -17.19% | 4.571% | -2.479% | -1.304%
| Z-Score (Absolute) |  | 0.95 | 0.4 | 2.84 | 1.2 | 0.11 | 0.11
| Analytical results for Pt in OREAS 241b (Indicative Value 8.55 ppb) | [32 empty] | Indicative
Pt | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.11 | 1.22 | 1.27 | 1.28 | 1.34 | [25 empty] | ppb
|  |  | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 01g | 01g | 25g | 0.5g | 0.5g | [25 empty] | 2
| 1 | 1 | 9 | 9 | < 5 | < 10 | 12.93 | < 10 | [25 empty] | 1
| 1 | 2 | 8 | 11 | < 5 | < 10 | 9.15 | 10 | [25 empty] | 1
| 1 | 3 | 8 | 8 | < 5 | < 10 | 16.62 | < 10 | [25 empty] | 16
| 1 | 4 | 8 | 6 | < 5 | < 10 | 14.67 | < 10 | [25 empty] | 8.5458
| 1 | 5 | 8 | 7 | < 5 | < 10 | 10.06 | < 10 | [25 empty] | 12
| 1 | 6 | 7 | 11 | < 5 | < 10 | 6.67 | < 10
| Mean |  | 8 | 8.6667 |  |  | 11.6833 | 10
| Median |  | 8 | 8.5 |  |  | 11.495 | 10
| Std Dev. |  | 0.63246 | 2.0656 |  |  | 3.7157
| Rel.Std.Dev. |  | 7.906% | 23.83% |  |  | 31.8%
| PDM3 |  | -6.387% | 1.414% |  |  | 36.71% | 17.02%
| Z-Score (Absolute) |  | 0.4 | 0.64 | 1.62 | 0.7 | 1.75 | 0.4
| Analytical results for Rb in OREAS 241b (Certified Value 7.91 ppm) | [32 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [10 empty] | 3
| 1 | 1 | 7.9 | 8.08 | 6.9605 | 7.7753 | 6.3 | 8.76 | 8.33 | 8.1 | 8.8 | 7.4 | 7.7 | 7.2 | 7.8 | 7.21 | 8.3 | 8.5 | 8.2 | 7.3 | 8 | 8.5 | 7.6 | [10 empty] | 1
| 1 | 2 | 7.5 | 8.12 | 6.9887 | 7.753 | 6.7 | 8.82 | 8.79 | 8.27 | 8.7 | 7.6 | 7.5 | 7.1 | 7.9 | 7.28 | 8.2 | 8 | 8.1 | 7.9 | 7.7 | 8.2 | 7.6 | [10 empty] | 30
| 1 | 3 | 7.8 | 8.03 | 6.9487 | 7.6994 | 6.5 | 8.86 | 8.85 | 8.11 | 8.8 | 7.6 | 7.6 | 7.1 | 7.8 | 7.5 | 8.3 | 8.5 | 8.4 | 7.5 | 8 | 8 | 7.8 | [10 empty] | 16
| 1 | 4 | 7.8 | 8.09 | 7.092 | 7.7402 | 6.3 | 8.73 | 8.74 | 8.03 | 9 | 7.4 | 7.7 | 7.2 | 7.7 | 7.5 | 8.7 | 8 | 7.8 | 7.8 | 7.9 | 8.1 | 7.5 | [10 empty] | 7.9145
| 1 | 5 | 7.8 | 8.09 | 6.9872 | 7.7722 | 6.5 | 8.67 | 8.84 | 7.63 | 9 | 7.5 | 7.7 | 7.1 | 7.6 | 7.17 | 8.7 | 8.5 | 8.6 | 7.6 | 8.1 | 8 | 7.7 | [10 empty] | 110
| 1 | 6 | 7.6 | 8.04 | 7.0689 | 7.6232 | 6.1 | 8.57 | 8.26 | 7.65 | 9 | 7.4 | 7.8 | 7.1 | 7.7 | 7.25 | 8.1 | 8.5 | 8.3 | 7.5 | 7.9 | 8 | 7.6
| Mean |  | 7.7333 | 8.075 | 7.0077 | 7.7272 | 6.4 | 8.735 | 8.635 | 7.965 | 8.8833 | 7.4833 | 7.6667 | 7.1333 | 7.75 | 7.3183 | 8.3833 | 8.3333 | 8.2333 | 7.6 | 7.9333 | 8.1333 | 7.6333
| Median |  | 7.8 | 8.085 | 6.988 | 7.7466 | 6.4 | 8.745 | 8.765 | 8.065 | 8.9 | 7.45 | 7.7 | 7.1 | 7.75 | 7.265 | 8.3 | 8.5 | 8.25 | 7.55 | 7.95 | 8.05 | 7.6
| Std Dev. |  | 0.15055 | 0.033912 | 0.058889 | 0.057912 | 0.20976 | 0.10483 | 0.26719 | 0.2638 | 0.13292 | 0.098319 | 0.10328 | 0.05164 | 0.10488 | 0.14552 | 0.25626 | 0.2582 | 0.27325 | 0.21909 | 0.13663 | 0.19664 | 0.10328
| Rel.Std.Dev. |  | 1.947% | 0.42% | 0.8404% | 0.7495% | 3.278% | 1.2% | 3.094% | 3.312% | 1.496% | 1.314% | 1.347% | 0.7239% | 1.353% | 1.988% | 3.057% | 3.098% | 3.319% | 2.883% | 1.722% | 2.418% | 1.353%
| PDM3 |  | -2.289% | 2.028% | -11.46% | -2.366% | -19.14% | 10.37% | 9.104% | 0.6381% | 12.24% | -5.448% | -3.131% | -9.87% | -2.078% | -7.533% | 5.924% | 5.292% | 4.029% | -3.974% | 0.238% | 2.765% | -3.552%
| Z-Score (Absolute) |  | 0.03 | 0.57 | 1.31 | 0.04 | 2.37 | 1.73 | 1.56 | 0.38 | 1.99 | 0.47 | 0.15 | 1.08 | 0 | 0.76 | 1.11 | 1.03 | 0.85 | 0.26 | 0.32 | 0.67 | 0.21
| Analytical results for Re in OREAS 241b (Certified Value 0.002 ppm) | [32 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 1.01 | 1.02 | 1.07 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [13 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [13 empty] | 3
|  |  | 0.25g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [13 empty] | 3
| 1 | 1 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.003 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.011 | < 0.002 | 0.003 | 0.002 | [13 empty] | 1
| 1 | 2 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.011 | < 0.002 | 0.002 | 0.002 | [13 empty] | 31
| 1 | 3 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.014 | < 0.002 | 0.002 | 0.001 | [13 empty] | 16
| 1 | 4 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | < 0.005 | 0.003 | 0.002 | 0.003 | 0.003 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.012 | < 0.002 | 0.002 | < 0.001 | [13 empty] | 0.00204
| 1 | 5 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.002 | 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.002 | 0.013 | < 0.002 | 0.002 | < 0.001 | [13 empty] | 111
| 1 | 6 | < 0.1 | 0.002 | 0.002 | < 0.05 | 0.003 | < 0.005 | 0.004 | 0.002 | 0.002 | 0.002 | < 5 | < 0.05 | < 0.05 | 0.001 | 0.01 | < 0.002 | 0.002 | < 0.001
| Mean |  |  | 0.002 | 0.002 |  | 0.0021667 | 0.005 | 0.0026667 | 0.002 | 0.0021667 | 0.0021667 |  |  |  | 0.0016667 | 0.011833 |  | 0.0021667 | 0.0016667
| Median |  |  | 0.002 | 0.002 |  | 0.002 | 0.005 | 0.0025 | 0.002 | 0.002 | 0.002 |  |  |  | 0.002 | 0.0115 |  | 0.002 | 0.002
| Std Dev. |  |  | 0 | 0 |  | 0.00040825 |  | 0.0008165 | 0 | 0.00040825 | 0.00040825 |  |  |  | 0.0005164 | 0.001472 |  | 0.00040825 | 0.00057735
| Rel.Std.Dev. |  |  | 0% | 0% |  | 18.84% |  | 30.62% | 0% | 18.84% | 18.84% |  |  |  | 30.98% | 12.44% |  | 18.84% | 34.64%
| PDM3 |  |  | -1.961% | -1.961% |  | 6.209% | 145.1% | 30.72% | -1.961% | 6.209% | 6.209% |  |  |  | -18.3% | 480.1% |  | 6.209% | -18.3%
| Z-Score (Absolute) |  | 51.61 | 0.18 | 0.18 | 24.63 | 0 | 0.81 | 0.54 | 0.18 | 0 | 0 | 2694.9 | 24.63 | 24.63 | 0.54 | 10.43 | 1.26 | 0 | 1.17
| Analytical results for S in OREAS 241b (Certified Value 0.496 wt.%) | [32 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.51 | 0.49 | 0.52702 | 0.51626 | 0.49618 | 0.474 | 0.49 | 0.51 | 0.49 | 0.48 | 0.49 | 0.52 | 0.49 | 0.48 | 0.51 | 0.43 | 0.45554 | 0.49 | 0.44 | 0.51 | 0.50161 | 0.4885 | 0.47 | 0.49 | 0.53 | 0.5 | 0.52 |  |  |  |  | 1
| 1 | 2 | 0.51 | 0.48 | 0.56805 | 0.50327 | 0.50843 | 0.48 | 0.495 | 0.49 | 0.5 | 0.48 | 0.49 | 0.5 | 0.48 | 0.49 | 0.5 | 0.43 | 0.46414 | 0.5 | 0.44 | 0.5 | 0.50157 | 0.496 | 0.47 | 0.49 | 0.51 | 0.49 | 0.52 |  |  |  |  | 16
| 1 | 3 | 0.51 | 0.48 | 0.53691 | 0.51336 | 0.50087 | 0.475 | 0.488 | 0.51 | 0.5 | 0.49 | 0.49 | 0.51 | 0.49 | 0.49 | 0.5 | 0.43 | 0.45036 | 0.5 | 0.43 | 0.5 | 0.50421 | 0.505 | 0.47 | 0.52 | 0.52 | 0.49 | 0.52 |  |  |  |  | 16
| 1 | 4 | 0.51 | 0.48 | 0.55958 | 0.51822 | 0.50957 | 0.477 | 0.491 | 0.49 | 0.48 | 0.5 | 0.49 | 0.51 | 0.49 | 0.49 | 0.5 | 0.43 | 0.4766 | 0.49 | 0.42 | 0.5 | 0.48932 | 0.4915 | 0.47 | 0.48 | 0.53 | 0.49 | 0.51 |  |  |  |  | 0.49617
| 1 | 5 | 0.53 | 0.49 | 0.52922 | 0.50815 | 0.50838 | 0.475 | 0.499 | 0.51 | 0.48 | 0.47 | 0.49 | 0.51 | 0.5 | 0.49 | 0.5 | 0.43 | 0.45046 | 0.49 | 0.45 | 0.5 | 0.48487 | 0.499 | 0.47 | 0.5 | 0.51 | 0.49 | 0.53 |  |  |  |  | 112
| 1 | 6 | 0.52 | 0.49 | 0.54371 | 0.5007 | 0.49724 | 0.475 | 0.493 | 0.51 | 0.46 | 0.48 | 0.48 | 0.5 | 0.49 | 0.49 | 0.5 | 0.43 | 0.45021 | 0.5 | 0.44 | 0.5 | 0.49161 | 0.505 | 0.47 | 0.47 | 0.52 | 0.48 | 0.52
| Mean |  | 0.515 | 0.485 | 0.54408 | 0.50999 | 0.50344 | 0.476 | 0.49267 | 0.50333 | 0.485 | 0.48333 | 0.48833 | 0.50833 | 0.49 | 0.48833 | 0.50167 | 0.43 | 0.45788 | 0.495 | 0.43667 | 0.50167 | 0.49553 | 0.4975 | 0.47 | 0.49167 | 0.52 | 0.49 | 0.52
| Median |  | 0.51 | 0.485 | 0.54031 | 0.51076 | 0.50462 | 0.475 | 0.492 | 0.51 | 0.485 | 0.48 | 0.49 | 0.51 | 0.49 | 0.49 | 0.5 | 0.43 | 0.453 | 0.495 | 0.44 | 0.5 | 0.49659 | 0.4975 | 0.47 | 0.49 | 0.52 | 0.49 | 0.52
| Std Dev. |  | 0.0083666 | 0.0054772 | 0.016604 | 0.0071176 | 0.0060776 | 0.0021909 | 0.0039328 | 0.010328 | 0.015166 | 0.010328 | 0.0040825 | 0.0075277 | 0.0063246 | 0.0040825 | 0.0040825 | 0 | 0.010636 | 0.0054772 | 0.010328 | 0.0040825 | 0.0079524 | 0.0068411 | 0 | 0.017224 | 0.0089443 | 0.0063246 | 0.0063246
| Rel.Std.Dev. |  | 1.625% | 1.129% | 3.052% | 1.396% | 1.207% | 0.4603% | 0.7983% | 2.052% | 3.127% | 2.137% | 0.836% | 1.481% | 1.291% | 0.836% | 0.8138% | 0% | 2.323% | 1.107% | 2.365% | 0.8138% | 1.605% | 1.375% | 0% | 3.503% | 1.72% | 1.291% | 1.216%
| PDM3 |  | 3.796% | -2.25% | 9.658% | 2.787% | 1.467% | -4.064% | -0.7051% | 1.445% | -2.25% | -2.586% | -1.578% | 2.452% | -1.243% | -1.578% | 1.109% | -13.34% | -7.716% | -0.2349% | -11.99% | 1.109% | -0.1276% | 0.269% | -5.274% | -0.9067% | 4.804% | -1.243% | 4.804%
| Z-Score (Absolute) |  | 1.61 | 0.55 | 3.71 | 1.25 | 0.78 | 1.2 | 0 | 0.77 | 0.55 | 0.67 | 0.31 | 1.13 | 0.19 | 0.31 | 0.65 | 4.53 | 2.51 | 0.17 | 4.05 | 0.65 | 0.21 | 0.35 | 1.64 | 0.07 | 1.97 | 0.19 | 1.97
| Analytical results for Sb in OREAS 241b (Certified Value 1.22 ppm) | [32 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 2
| 1 | 1 | 1.67 | 1.45 | 4.56 | 1.4037 | 1.2 | 1.53 | 0.36 | 0.62 | < 5 | 1.56 | 1.24 | 1.57 | 1.45 | 96.573 | 1.31 | < 5 | < 5 | 1.27 | 1.11 | 0.84 | 5.65 | 1 | 1.6037 | 1.38 | 0.306 | 0.8 | 0.74 | 0.72 | 1.06 |  |  | 1
| 1 | 2 | 1.71 | 1.42 | 4.97 | 1.445 | 1.28 | 1.56 | 0.31 | 0.63 | < 5 | 1.64 | 1.24 | 1.8 | 1.41 | 104.089 | 1.26 | < 5 | < 5 | 1.26 | 1.08 | 0.78 | 6.62 | 1 | 1.6406 | 1.38 | 0.314 | 0.83 | 0.67 | 0.7 | 0.96 |  |  | 32
| 1 | 3 | 1.69 | 1.4 | 5.9 | 1.2959 | 1.3 | 1.5 | 0.37 | 0.59 | < 5 | 1.52 | 1.36 | 1.57 | 1.41 | 99.528 | 1.27 | < 5 | < 5 | 1.32 | 1.02 | 0.79 | 6.73 | 1 | 1.6499 | 1.41 | 0.282 | 0.79 | 0.69 | 0.71 | 0.96 |  |  | 16
| 1 | 4 | 1.69 | 1.41 | 6.14 | 1.3817 | 1.26 | 1.51 | 0.39 | 0.66 | < 5 | 1.5 | 1.32 | 1.54 | 1.48 | 91.448 | 1.27 | < 5 | < 5 | 1.3 | 1.12 | 0.82 | 7.2 | 1 | 1.6028 | 1.38 | 0.325 | 0.77 | 0.7 | 0.73 | 0.94 |  |  | 1.2173
| 1 | 5 | 1.72 | 1.43 | 4.48 | 1.3991 | 1.18 | 1.49 | 0.35 | 0.63 | < 5 | 1.46 | 1.3 | 1.6 | 1.38 | 96.365 | 1.26 | < 5 | < 5 | 1.23 | 0.97 | 0.83 | 5.98 | 1 | 1.5851 | 1.44 | 0.331 | 0.79 | 0.72 | 0.75 | 0.92 |  |  | 113
| 1 | 6 | 1.71 | 1.43 | 5.8 | 1.344 | 1.2 | 1.53 | 0.42 | 0.69 | < 5 | 1.56 | 1.34 | 1.62 | 1.48 | 103.243 | 1.26 | < 5 | < 5 | 1.32 | 1.01 | 0.8 | 8.07 | 1 | 1.5756 | 1.42 | 0.317 | 0.8 | 0.7 | 0.73 | 0.89
| Mean |  | 1.6983 | 1.4233 | 5.3083 | 1.3782 | 1.2367 | 1.52 | 0.36667 | 0.63667 |  | 1.54 | 1.3 | 1.6167 | 1.435 | 98.541 | 1.2717 |  |  | 1.2833 | 1.0517 | 0.81 | 6.7083 | 1 | 1.6096 | 1.4017 | 0.3125 | 0.79667 | 0.70333 | 0.72333 | 0.955
| Median |  | 1.7 | 1.425 | 5.385 | 1.3904 | 1.23 | 1.52 | 0.365 | 0.63 |  | 1.54 | 1.31 | 1.585 | 1.43 | 98.0505 | 1.265 |  |  | 1.285 | 1.05 | 0.81 | 6.675 | 1 | 1.6033 | 1.395 | 0.3155 | 0.795 | 0.7 | 0.725 | 0.95
| Std Dev. |  | 0.018348 | 0.017512 | 0.7272 | 0.051969 | 0.049666 | 0.025298 | 0.037238 | 0.034448 |  | 0.061968 | 0.050596 | 0.09395 | 0.041352 | 4.7494 | 0.019408 |  |  | 0.036148 | 0.06047 | 0.023664 | 0.86657 | 0 | 0.02974 | 0.025626 | 0.017283 | 0.019664 | 0.024221 | 0.017512 | 0.057879
| Rel.Std.Dev. |  | 1.08% | 1.23% | 13.7% | 3.771% | 4.016% | 1.664% | 10.16% | 5.411% |  | 4.024% | 3.892% | 5.811% | 2.882% | 4.82% | 1.526% |  |  | 2.817% | 5.75% | 2.922% | 12.92% | 0% | 1.848% | 1.828% | 5.531% | 2.468% | 3.444% | 2.421% | 6.061%
| PDM3 |  | 39.51% | 16.92% | 336.1% | 13.22% | 1.587% | 24.86% | -69.88% | -47.7% |  | 26.51% | 6.79% | 32.8% | 17.88% | 7995% | 4.463% |  |  | 5.421% | -13.61% | -33.46% | 451.1% | -17.85% | 32.22% | 15.14% | -74.33% | -34.56% | -42.22% | -40.58% | -21.55%
| Z-Score (Absolute) |  | 0.54 | 0.06 | 6.83 | 0.02 | 0.27 | 0.23 | 1.78 | 1.31 | 1.94 | 0.26 | 0.16 | 0.4 | 0.08 | 169.45 | 0.21 | 1.94 | 1.94 | 0.19 | 0.59 | 1.01 | 9.28 | NA | 0.38 | 0.02 | 1.88 | 1.03 | 1.2 | 1.16 | 0.76
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sc in OREAS 241b (Certified Value 5.31 ppm) | [32 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 5 | 5.68 | 5.466 | 6.74 | 5.5852 | 6.7 | 5 | 7 | 6.5 | 6 | 4.9 | 5.9 | 4.7 | 5.4 | 5.5 | 5 | 6 | 4.5 | 5.4 | 4 | 5.8 | 1.9 | 5.3 | 5.7 | 4.3 | [6 empty] | 1
| 1 | 2 | 4.8 | 5.79 | 5.379 | 6.97 | 5.7291 | 6.6 | 5 | 7 | 6.6 | 6 | 5 | 5.9 | 4.7 | 5.4 | 5.6 | 5 | 6 | 4.9 | 5.1 | 4 | 5.8 | 2.1 | 5.6 | 5.4 | 4.3 | [6 empty] | 33
| 1 | 3 | 4.8 | 5.72 | 4.898 | 6.7 | 5.6053 | 7.3 | 5 | 7 | 6.7 | 6 | 5.1 | 6 | 4.7 | 5.4 | 5.6 | 5 | 6 | 4.7 | 5.3 | 5 | 5.7 | 2.2 | 5.5 | 5.2 | 4.6 | [6 empty] | 16
| 1 | 4 | 4.8 | 5.71 | 5.061 | 7.48 | 5.6058 | 6.8 | 5 | 7 | 6.6 | 6 | 5.1 | 5.8 | 4.7 | 5.4 | 5.5 | 5 | 6 | 4.3 | 5.3 | 4 | 5.6 | 2.1 | 5.6 | 5.4 | 4.3 | [6 empty] | 5.3109
| 1 | 5 | 4.7 | 5.73 | 4.94 | 7.09 | 5.7512 | 6.6 | 5 | 7 | 6.1 | 6 | 5 | 5.8 | 4.6 | 5.3 | 5.4 | 4 | 6 | 4.8 | 5.3 | 5 | 5.7 | 2.4 | 5.2 | 5.3 | 4.4 | [6 empty] | 114
| 1 | 6 | 4.6 | 5.65 | 5.088 | 6.97 | 5.5403 | 6.8 | 5 | 7 | 6.3 | 6 | 5.2 | 5.8 | 4.7 | 5.3 | 5.4 | 4 | 6 | 4.7 | 5.1 | 5 | 5.7 | 1.8 | 5.5 | 5.3 | 4.3
| Mean |  | 4.7833 | 5.7133 | 5.1387 | 6.9917 | 5.6362 | 6.8 | 5 | 7 | 6.4667 | 6 | 5.05 | 5.8667 | 4.6833 | 5.3667 | 5.5 | 4.6667 | 6 | 4.65 | 5.25 | 4.5 | 5.7167 | 2.0833 | 5.45 | 5.3833 | 4.3667
| Median |  | 4.8 | 5.715 | 5.0745 | 6.97 | 5.6055 | 6.75 | 5 | 7 | 6.55 | 6 | 5.05 | 5.85 | 4.7 | 5.4 | 5.5 | 5 | 6 | 4.7 | 5.3 | 4.5 | 5.7 | 2.1 | 5.5 | 5.35 | 4.3
| Std Dev. |  | 0.13292 | 0.04761 | 0.23276 | 0.28195 | 0.084323 | 0.26077 | 0 | 0 | 0.22509 | 0 | 0.10488 | 0.08165 | 0.040825 | 0.05164 | 0.089443 | 0.5164 | 0 | 0.21679 | 0.12247 | 0.54772 | 0.075277 | 0.2137 | 0.16432 | 0.17224 | 0.12111
| Rel.Std.Dev. |  | 2.779% | 0.8333% | 4.53% | 4.033% | 1.496% | 3.835% | 0% | 0% | 3.481% | 0% | 2.077% | 1.392% | 0.8717% | 0.9622% | 1.626% | 11.07% | 0% | 4.662% | 2.333% | 12.17% | 1.317% | 10.26% | 3.015% | 3.2% | 2.773%
| PDM3 |  | -9.934% | 7.577% | -3.243% | 31.65% | 6.124% | 28.04% | -5.854% | 31.8% | 21.76% | 12.97% | -4.913% | 10.46% | -11.82% | 1.05% | 3.56% | -12.13% | 12.97% | -12.44% | -1.147% | -15.27% | 7.64% | -60.77% | 2.619% | 1.363% | -17.78%
| Z-Score (Absolute) |  | 1.21 | 0.67 | 0.49 | 3.25 | 0.51 | 2.87 | NA | NA | 2.19 | NA | 0.67 | 0.98 | 1.42 | 0.03 | 0.24 | NA | NA | 1.48 | 0.27 | NA | 0.67 | 6.68 | 0.13 | 0 | 2.06
Results from laboratories 1.08, 1.10, 1.12, 1.19, 1.22 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 241b (Indicative Value 0.71 ppm) | [32 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 2
| 1 | 1 | < 1 | 0.52 | < 1 | < 1.5 | 0.7 | 8 | 1 | < 1 | < 5 | 2.3 | 0.9 | 0.6 | 0.9 | 1 | < 5 | 6.8173 | < 1 | 0.5 | 2.57 | < 1 | 0.42789 | 0.7 | 0.424 | 0.9 | < 1 | 0.6 | 0.5 |  |  |  |  | 1
| 1 | 2 | < 1 | 0.43 | < 1 | < 1.5 | 0.7 | < 1 | 1 | < 1 | < 5 | 2 | 1 | 0.5 | 0.7 | 0.7 | < 5 | < 5 | < 1 | 0.5 | 3.33 | < 1 | 0.57909 | 0.8 | 0.505 | 1 | < 1 | 0.5 | 0.4 |  |  |  |  | 7
| 1 | 3 | < 1 | 0.48 | < 1 | < 1.5 | 1 | 3 | 1 | < 1 | < 5 | 2.7 | 0.9 | 0.5 | 0.8 | 0.9 | < 5 | 5.175 | < 1 | 0.6 | 3.84 | < 1 | 0.55219 | 0.8 | 0.523 | 1.3 | < 1 | 0.5 | 0.5 |  |  |  |  | 16
| 1 | 4 | < 1 | 0.46 | < 1 | < 1.5 | 0.8 | 5 | 1 | < 1 | < 5 | 1.9 | 0.9 | 0.6 | 1.1 | 0.8 | < 5 | < 5 | < 1 | 0.5 | 3.84 | < 1 | 0.53696 | 0.8 | 0.453 | 1.5 | < 1 | 0.6 | 0.5 |  |  |  |  | 0.7056
| 1 | 5 | < 1 | 0.49 | < 1 | < 1.5 | 0.7 | < 1 | 1 | < 1 | < 5 | 2.7 | 0.4 | 0.6 | 0.8 | 1 | < 5 | < 5 | < 1 | 0.5 | < 0.1 | < 1 | 0.53413 | 0.8 | 0.502 | 1.6 | < 1 | 0.5 | 0.5 |  |  |  |  | 13
| 1 | 6 | < 1 | 0.46 | < 1 | < 1.5 | 0.7 | < 1 | 1 | < 1 | < 5 | 2.2 | 0.9 | 0.6 | 1 | 1 | < 5 | < 5 | < 1 | 0.4 | 2.28 | < 1 | 0.51618 | 0.7 | 0.421 | 1.2 | < 1 | 0.5 | 0.6
| Mean |  |  | 0.47333 |  |  | 0.76667 | 5.3333 | 1 |  |  | 2.3 | 0.83333 | 0.56667 | 0.88333 | 0.9 |  | 5.9962 |  | 0.5 | 3.172 |  | 0.52441 | 0.76667 | 0.47133 | 1.25 |  | 0.53333 | 0.5
| Median |  |  | 0.47 |  |  | 0.7 | 5 | 1 |  |  | 2.25 | 0.9 | 0.6 | 0.85 | 0.95 |  | 5.9962 |  | 0.5 | 3.33 |  | 0.53554 | 0.8 | 0.4775 | 1.25 |  | 0.5 | 0.5
| Std Dev. |  |  | 0.030768 |  |  | 0.12111 | 2.5166 | 0 |  |  | 0.34059 | 0.21602 | 0.05164 | 0.1472 | 0.12649 |  | 1.1613 |  | 0.063246 | 0.72033 |  | 0.051764 | 0.05164 | 0.044392 | 0.27386 |  | 0.05164 | 0.063246
| Rel.Std.Dev. |  |  | 6.5% |  |  | 15.8% | 47.19% | 0% |  |  | 14.81% | 25.92% | 9.113% | 16.66% | 14.05% |  | 19.37% |  | 12.65% | 22.71% |  | 9.871% | 6.736% | 9.418% | 21.91% |  | 9.682% | 12.65%
| PDM3 |  |  | -32.92% |  |  | 8.655% | 655.9% | 41.72% |  |  | 226% | 18.1% | -19.69% | 25.19% | 27.55% |  | 749.8% |  | -29.14% | 349.5% |  | -25.68% | 8.655% | -33.2% | 77.15% |  | -24.41% | -29.14%
| Z-Score (Absolute) |  | 0.67 | 0.75 | 0.67 | 0 | 0.04 | 5.84 | 0.67 | 0.67 | 4.72 | 4.18 | 0.22 | 0.49 | 0.36 | 0.4 | 4.72 | 7.86 | 0.67 | 0.67 | 5.13 | 0.67 | 0.61 | 0.04 | 0.75 | 1.35 | 0.67 | 0.58 | 0.67
| Analytical results for Si in OREAS 241b (Indicative Value 0.11 wt.%) | [32 empty] | Indicative
Si | Round | Replicate | Lab | Lab | [29 empty] | 1
| No | No | 1.25 | 1.27 | [29 empty] | wt.%
|  |  | AR*OES/MS | AR*MS | [29 empty] | 3
|  |  | 0.5g | 25g | [29 empty] | 3
| 1 | 1 | 0.1692 | 0.0675 | [29 empty] | 1
| 1 | 2 | 0.1376 | 0.0815 | [29 empty] | 5
| 1 | 3 | 0.12699 | 0.0765 | [29 empty] | 16
| 1 | 4 | 0.18282 | 0.077 | [29 empty] | 0.11038
| 1 | 5 | 0.113 | 0.091 | [29 empty] | 11
| 1 | 6 | 0.11351 | 0.088
| Mean |  | 0.14052 | 0.08025
| Median |  | 0.13229 | 0.07925
| Std Dev. |  | 0.029293 | 0.0085367
| Rel.Std.Dev. |  | 20.85% | 10.64%
| PDM3 |  | 27.3% | -27.3%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sm in OREAS 241b (Certified Value 1.6 ppm) | [32 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [25 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | [25 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | [25 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [25 empty] | 2
| 1 | 1 | 1.662 | 1.607 | 1.3659 | 1.7 | 1.79 | 1.4 | [25 empty] | 1
| 1 | 2 | 1.672 | 1.606 | 1.3449 | 1.8 | 1.8 | 1.4 | [25 empty] | 18
| 1 | 3 | 1.663 | 1.505 | 1.3869 | 1.6 | 1.81 | 1.3 | [25 empty] | 16
| 1 | 4 | 1.677 | 1.674 | 1.3162 | 1.7 | 1.81 | 1.3 | [25 empty] | 1.5959
| 1 | 5 | 1.675 | 1.595 | 1.3412 | 1.8 | 1.88 | 1.4 | [25 empty] | 115
| 1 | 6 | 1.675 | 1.611 | 1.3726 | 1.8 | 1.8 | 1.5
| Mean |  | 1.6707 | 1.5997 | 1.3546 | 1.7333 | 1.815 | 1.3833
| Median |  | 1.6735 | 1.6065 | 1.3554 | 1.75 | 1.805 | 1.4
| Std Dev. |  | 0.006532 | 0.05428 | 0.025471 | 0.08165 | 0.032711 | 0.075277
| Rel.Std.Dev. |  | 0.391% | 3.393% | 1.88% | 4.711% | 1.802% | 5.442%
| PDM3 |  | 4.683% | 0.2346% | -15.12% | 8.61% | 13.73% | -13.32%
| Z-Score (Absolute) |  | 0.17 | 0.17 | 1.36 | 0.48 | 0.87 | 1.22
| Analytical results for Sn in OREAS 241b (Certified Value 0.73 ppm) | [32 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  |  | 3
| 1 | 1 | 0.8 | 0.75 | < 1 | 0.52838 | 0.67691 | 0.79 | 0.62 | 0.8 | 0.77 | < 10 | 0.6 | 0.7 | 0.7 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.7 | 40.5 | 0.5 | 0.75383 | 0.7 | 0.8 | 0.8 | 0.8 | 0.5 |  |  |  |  | 1
| 1 | 2 | 0.8 | 0.72 | < 1 | 0.54885 | 0.64512 | 0.8 | 0.62 | 0.9 | 0.79 | < 10 | 0.7 | 0.7 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.8 | 0.7 | 40.53 | 0.5 | 0.77056 | 0.7 | 0.7 | 0.7 | 0.8 | 0.5 |  |  |  |  | 19
| 1 | 3 | 0.8 | 0.74 | < 1 | 0.55476 | 0.70456 | 0.81 | 0.61 | 0.8 | 0.74 | < 10 | 0.6 | 0.7 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.7 | 0.7 | 43.27 | 0.5 | 0.76531 | 0.7 | 0.8 | 0.9 | 0.7 | 0.5 |  |  |  |  | 16
| 1 | 4 | 0.8 | 0.71 | < 1 | 0.55343 | 0.69556 | 0.78 | 0.62 | 0.9 | 0.76 | < 10 | 0.6 | 0.7 | 0.7 | 0.7 | 0.7 | < 10 | < 5 | 0.7 | 0.8 | 45.74 | 0.5 | 0.7518 | 0.7 | 0.7 | 0.8 | 0.8 | 0.5 |  |  |  |  | 0.73396
| 1 | 5 | 0.8 | 0.72 | < 1 | 0.54991 | 0.64016 | 0.8 | 0.62 | 0.9 | 0.78 | < 10 | 0.7 | 0.7 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.7 | 0.8 | 43.22 | 0.5 | 0.76188 | 0.7 | 0.7 | 0.8 | 0.7 | 0.5 |  |  |  |  | 116
| 1 | 6 | 0.8 | 0.72 | < 1 | 0.53342 | 0.66441 | 0.73 | 0.66 | 0.9 | 0.71 | < 10 | 0.7 | 0.7 | 0.8 | 0.7 | 0.7 | < 10 | < 5 | 0.7 | 0.7 | 44.78 | 0.5 | 0.7468 | 0.7 | 0.6 | 0.8 | 0.7 | 0.4
| Mean |  | 0.8 | 0.72667 |  | 0.54479 | 0.67112 | 0.785 | 0.625 | 0.86667 | 0.75833 |  | 0.65 | 0.7 | 0.76667 | 0.7 | 0.7 |  |  | 0.73333 | 0.73333 | 43.0067 | 0.5 | 0.75836 | 0.7 | 0.71667 | 0.8 | 0.75 | 0.48333
| Median |  | 0.8 | 0.72 |  | 0.54938 | 0.67066 | 0.795 | 0.62 | 0.9 | 0.765 |  | 0.65 | 0.7 | 0.8 | 0.7 | 0.7 |  |  | 0.7 | 0.7 | 43.245 | 0.5 | 0.75785 | 0.7 | 0.7 | 0.8 | 0.75 | 0.5
| Std Dev. |  | 1.2162e-16 | 0.015055 |  | 0.011093 | 0.026186 | 0.02881 | 0.017607 | 0.05164 | 0.029269 |  | 0.054772 | 1.2162e-16 | 0.05164 | 1.2162e-16 | 1.2162e-16 |  |  | 0.05164 | 0.05164 | 2.1516 | 0 | 0.0090191 | 1.2162e-16 | 0.075277 | 0.063246 | 0.054772 | 0.040825
| Rel.Std.Dev. |  | 1.52e-14% | 2.072% |  | 2.036% | 3.902% | 3.67% | 2.817% | 5.958% | 3.86% |  | 8.427% | 1.737e-14% | 6.736% | 1.737e-14% | 1.737e-14% |  |  | 7.042% | 7.042% | 5.003% | 0% | 1.189% | 1.737e-14% | 10.5% | 7.906% | 7.303% | 8.447%
| PDM3 |  | 8.998% | -0.9931% |  | -25.77% | -8.561% | 6.955% | -14.84% | 18.08% | 3.321% |  | -11.44% | -4.626% | 4.457% | -4.626% | -4.626% |  |  | -0.08474% | -0.08474% | 5760% | -31.88% | 3.326% | -4.626% | -2.356% | 8.998% | 2.186% | -34.15%
| Z-Score (Absolute) |  | 0.72 | 0.07 | 2.53 | 2.04 | 0.67 | 0.56 | 1.17 | 1.45 | 0.27 | 46.25 | 0.9 | 0.36 | 0.36 | 0.36 | 0.36 | 46.25 | 19.15 | 0 | 0 | 458.19 | 2.53 | 0.27 | 0.36 | 0.18 | 0.72 | 0.18 | 2.71
| Analytical results for Sr in OREAS 241b (Certified Value 37.5 ppm) | [32 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  |  | 1
| 1 | 1 | 28.3 | 39.55 | 165.01 | 41.7315 | 37.6133 | 50.5 | 44 | 45.6 | 45.66 | 41 | 39 | 37.1 | 39.3 | 35.7 | 38.7 | 36 | 41.8533 | 32.7 | 33.1 | 39.1 | 33.91 | 39.8412 | 29 | 39 | 32.7 | 37.4 | 38.5 | 31.1 |  |  |  | 1
| 1 | 2 | 29 | 38.93 | 162.92 | 43.5165 | 38.5242 | 53.4 | 45 | 44.1 | 46.56 | 40 | 39 | 36.9 | 41.1 | 37 | 39.1 | 37 | 44.102 | 32.1 | 35.3 | 38.6 | 35.21 | 38.781 | 29 | 39.4 | 34.3 | 35.3 | 37.3 | 30.9 |  |  |  | 20
| 1 | 3 | 24.4 | 38.71 | 147.38 | 46.2787 | 37.5127 | 51.4 | 44 | 42.9 | 44.68 | 40 | 39 | 37.2 | 39.2 | 35.2 | 39 | 36 | 42.9733 | 35.8 | 34 | 39.2 | 33.73 | 37.9998 | 30 | 39.3 | 34.7 | 36.5 | 36.8 | 31.7 |  |  |  | 16
| 1 | 4 | 24.5 | 38.77 | 162.26 | 46.2427 | 37.6528 | 48.7 | 44 | 44 | 44.64 | 42 | 39 | 36.4 | 38.6 | 36.5 | 38.6 | 36 | 41.988 | 35.7 | 31.6 | 39 | 33.73 | 35.4144 | 28 | 37.8 | 34.6 | 36.3 | 37.5 | 30.6 |  |  |  | 37.4521
| 1 | 5 | 25.6 | 39.32 | 143.88 | 42.7216 | 38.1772 | 50.2 | 45 | 44.6 | 42.55 | 41 | 40 | 37.6 | 39.6 | 36.4 | 37.1 | 35 | 39.8997 | 30.7 | 35.1 | 39.6 | 32.82 | 33.5544 | 30 | 39.8 | 34.4 | 36.3 | 36 | 31.1 |  |  |  | 117
| 1 | 6 | 26.1 | 38.53 | 179.23 | 42.5916 | 37.0914 | 49.4 | 45 | 47.1 | 43.61 | 42 | 40 | 37.4 | 39.1 | 35.5 | 37.7 | 35 | 40.2514 | 34.1 | 33.8 | 38.7 | 31.98 | 56.9532 | 31 | 39.5 | 31.6 | 35.1 | 36.2 | 30.9
| Mean |  | 26.3167 | 38.9683 | 160.1133 | 43.8471 | 37.7619 | 50.6 | 44.5 | 44.7167 | 44.6167 | 41 | 39.3333 | 37.1 | 39.4833 | 36.05 | 38.3667 | 35.8333 | 41.8446 | 33.5167 | 33.8167 | 39.0333 | 33.5633 | 40.424 | 29.5 | 39.1333 | 33.7167 | 36.15 | 37.05 | 31.05
| Median |  | 25.85 | 38.85 | 162.59 | 43.119 | 37.6331 | 50.35 | 44.5 | 44.35 | 44.66 | 41 | 39 | 37.15 | 39.25 | 36.05 | 38.65 | 36 | 41.9207 | 33.4 | 33.9 | 39.05 | 33.73 | 38.3904 | 29.5 | 39.35 | 34.35 | 36.3 | 37.05 | 31
| Std Dev. |  | 1.9323 | 0.3903 | 12.8652 | 1.9534 | 0.50977 | 1.6553 | 0.54772 | 1.4607 | 1.4251 | 0.89443 | 0.5164 | 0.41952 | 0.85654 | 0.6892 | 0.79415 | 0.75277 | 1.5944 | 2.0459 | 1.3644 | 0.36148 | 1.0906 | 8.4192 | 1.0488 | 0.70333 | 1.2703 | 0.84321 | 0.9225 | 0.36742
| Rel.Std.Dev. |  | 7.342% | 1.002% | 8.035% | 4.455% | 1.35% | 3.271% | 1.231% | 3.267% | 3.194% | 2.182% | 1.313% | 1.131% | 2.169% | 1.912% | 2.07% | 2.101% | 3.81% | 6.104% | 4.035% | 0.9261% | 3.249% | 20.83% | 3.555% | 1.797% | 3.768% | 2.333% | 2.49% | 1.183%
| PDM3 |  | -29.73% | 4.049% | 327.5% | 17.08% | 0.8274% | 35.11% | 18.82% | 19.4% | 19.13% | 9.473% | 5.023% | -0.94% | 5.424% | -3.744% | 2.442% | -4.322% | 11.73% | -10.51% | -9.707% | 4.222% | -10.38% | 7.935% | -21.23% | 4.489% | -9.974% | -3.477% | -1.074% | -17.09%
| Z-Score (Absolute) |  | 2.77 | 0.07 | 27.25 | 1.16 | 0.2 | 2.68 | 1.31 | 1.36 | 1.33 | 0.52 | 0.15 | 0.35 | 0.18 | 0.59 | 0.07 | 0.64 | 0.71 | 1.16 | 1.09 | 0.08 | 1.15 | 0.39 | 2.06 | 0.1 | 1.11 | 0.56 | 0.36 | 1.71
| Analytical results for Ta in OREAS 241b (Certified Value < 0.01 ppm) | [32 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.22 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 01g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [11 empty] | 3
| 1 | 1 | < 0.05 | < 0.01 | 0.01369 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.16 | 0.09 | < 0.01 | < 0.01 | [11 empty] | 1
| 1 | 2 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.03 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.14 | 0.07 | < 0.01 | < 0.01 | [11 empty] | 21
| 1 | 3 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.29 | 0.08 | < 0.01 | < 0.01 | [11 empty] | 16
| 1 | 4 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | 0.01 | 0.01 | < 0.01 | 0.01 | < 0.05 | < 0.5 | 0.01 | 0.15 | 0.06 | < 0.01 | < 0.01 | [11 empty] | < 0.01
| 1 | 5 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.16 | 0.06 | < 0.01 | < 0.01 | [11 empty] | 118
| 1 | 6 | < 0.05 | < 0.01 | < 0.01 | < 0.05 | 0.04 | < 0.05 | < 0.01 | 5 | < 0.1 | < 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.5 | 0.01 | 0.22 | 0.11 | < 0.01 | < 0.01
| Mean |  |  |  | 0.01369 |  | 0.036667 |  |  | 5 |  | 0.01 | 0.01 |  | 0.01 |  |  | 0.01 | 0.18667 | 0.078333
| Median |  |  |  | 0.01369 |  | 0.04 |  |  | 5 |  | 0.01 | 0.01 |  | 0.01 |  |  | 0.01 | 0.16 | 0.075
| Std Dev. | [5 empty] | 0.005164 |  |  | 0 |  | 0 | 0 |  |  |  |  | 0 | 0.05785 | 0.019408
| Rel.Std.Dev. | [5 empty] | 14.08% |  |  | 0% |  | 0% | 0% |  |  |  |  | 0% | 30.99% | 24.78%
PDM3
| Z-Score (Absolute) |  | 0.4 | 0.67 | 0.6 | 0.4 | 1.03 | 0.4 | 0.67 | 268.78 | 1.75 | 0.49 | 0.4 | 0.67 | 0.63 | 0.4 | 12.54 | 0.4 | 9.13 | 3.28 | 0.67 | 0.67
| Analytical results for Tb in OREAS 241b (Certified Value 0.31 ppm) | [32 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 3
| 1 | 1 | 0.32 | 0.332 | 0.311 | 0.29869 | 0.4 | 0.3 | 0.3 | 0.34 | [23 empty] | 1
| 1 | 2 | 0.32 | 0.324 | 0.323 | 0.30236 | 0.4 | 0.31 | 0.3 | 0.32 | [23 empty] | 22
| 1 | 3 | 0.31 | 0.316 | 0.301 | 0.30183 | 0.4 | 0.32 | 0.25 | 0.33 | [23 empty] | 16
| 1 | 4 | 0.32 | 0.323 | 0.32 | 0.30177 | 0.4 | 0.31 | 0.25 | 0.34 | [23 empty] | 0.31475
| 1 | 5 | 0.31 | 0.323 | 0.321 | 0.29684 | 0.4 | 0.31 | 0.3 | 0.34 | [23 empty] | 119
| 1 | 6 | 0.31 | 0.332 | 0.319 | 0.30516 | 0.4 | 0.32 | 0.3 | 0.32
| Mean |  | 0.315 | 0.325 | 0.31583 | 0.30111 | 0.4 | 0.31167 | 0.28333 | 0.33167
| Median |  | 0.315 | 0.3235 | 0.3195 | 0.3018 | 0.4 | 0.31 | 0.3 | 0.335
| Std Dev. |  | 0.0054772 | 0.0061319 | 0.0083526 | 0.002934 | 6.0809e-17 | 0.0075277 | 0.02582 | 0.0098319
| Rel.Std.Dev. |  | 1.739% | 1.887% | 2.645% | 0.9744% | 1.52e-14% | 2.415% | 9.113% | 2.964%
| PDM3 |  | 0.07977% | 3.257% | 0.3445% | -4.334% | 27.09% | -0.9793% | -9.981% | 5.375%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.06 | 0.94 | NA | 0.22 | 2.14 | 1.12
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 241b (Certified Value 0.13 ppm) | [32 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 15g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 0.11 | 0.13 | 0.16351 | 0.08 | 0.1 | 0.08 | 0.1 | < 5 | 0.1 | 0.14 | 0.14 | 0.12 | 0.14 | < 5 | 0.13 | 0.12 | 0.1 | 0.10395 | 0.14 | 0.18 | 0.1 | 0.15 | 0.11 | 0.18 | [7 empty] | 1
| 1 | 2 | 0.13 | 0.13 | 0.13521 | 0.11 | 0.11 | 0.09 | 0.1 | < 5 | 0.2 | 0.12 | 0.13 | 0.12 | 0.14 | < 5 | 0.12 | 0.12 | 0.1 | 0.15897 | 0.12 | 0.154 | 0.03 | 0.11 | 0.12 | 0.16 | [7 empty] | 23
| 1 | 3 | 0.1 | 0.12 | 0.1444 | 0.17 | 0.14 | 0.06 | 0.1 | < 5 | 0.2 | 0.14 | 0.13 | 0.11 | 0.13 | < 5 | 0.14 | 0.1 | 0.1 | 0.22149 | 0.13 | 0.129 | 0.07 | 0.14 | 0.13 | 0.16 | [7 empty] | 16
| 1 | 4 | 0.12 | 0.13 | 0.15493 | 0.12 | 0.12 | 0.08 | 0.1 | < 5 | 0.1 | 0.13 | 0.14 | 0.13 | 0.13 | < 5 | 0.11 | 0.12 | 0.1 | 0.1373 | 0.13 | 0.18 | 0.04 | 0.12 | 0.12 | 0.08 | [7 empty] | 0.12598
| 1 | 5 | 0.14 | 0.14 | 0.15953 | 0.13 | 0.12 | 0.09 | 0.1 | < 5 | 0.1 | 0.11 | 0.14 | 0.12 | 0.13 | < 5 | 0.1 | 0.1 | 0.1 | 0.16154 | 0.13 | < 0.03 | < 0.01 | 0.14 | 0.13 | 0.11 | [7 empty] | 120
| 1 | 6 | 0.16 | 0.12 | 0.16252 | 0.11 | 0.13 | 0.09 | 0.1 | < 5 | 0.1 | 0.11 | 0.13 | 0.11 | 0.12 | < 5 | 0.12 | 0.1 | 0.1 | 0.12027 | 0.14 | 0.18 | 0.09 | 0.13 | 0.11 | 0.12
| Mean |  | 0.12667 | 0.12833 | 0.15335 | 0.12 | 0.12 | 0.081667 | 0.1 |  | 0.13333 | 0.125 | 0.135 | 0.11833 | 0.13167 |  | 0.12 | 0.11 | 0.1 | 0.15059 | 0.13167 | 0.1646 | 0.066 | 0.13167 | 0.12 | 0.135
| Median |  | 0.125 | 0.13 | 0.15723 | 0.115 | 0.12 | 0.085 | 0.1 |  | 0.1 | 0.125 | 0.135 | 0.12 | 0.13 |  | 0.12 | 0.11 | 0.1 | 0.14814 | 0.13 | 0.18 | 0.07 | 0.135 | 0.12 | 0.14
| Std Dev. |  | 0.021602 | 0.0075277 | 0.011289 | 0.029665 | 0.014142 | 0.01169 | 1.5202e-17 |  | 0.05164 | 0.013784 | 0.0054772 | 0.0075277 | 0.0075277 |  | 0.014142 | 0.010954 | 1.5202e-17 | 0.041202 | 0.0075277 | 0.022865 | 0.030496 | 0.01472 | 0.0089443 | 0.037815
| Rel.Std.Dev. |  | 17.05% | 5.866% | 7.362% | 24.72% | 11.79% | 14.31% | 1.52e-14% |  | 38.73% | 11.03% | 4.057% | 6.361% | 5.717% |  | 11.79% | 9.959% | 1.52e-14% | 27.36% | 5.717% | 13.89% | 46.21% | 11.18% | 7.454% | 28.01%
| PDM3 |  | 0.543% | 1.866% | 21.72% | -4.749% | -4.749% | -35.18% | -20.62% |  | 5.835% | -0.78% | 7.158% | -6.072% | 4.512% |  | -4.749% | -12.69% | -20.62% | 19.53% | 4.512% | 30.65% | -47.61% | 4.512% | -4.749% | 7.158%
| Z-Score (Absolute) |  | 0.13 | 0 | 2.02 | 0.67 | 0.67 | 3.78 | NA | 191.91 | NA | 0.27 | 0.54 | 0.81 | 0.27 | 191.91 | 0.67 | 1.48 | NA | 1.8 | 0.27 | 0.92 | 5.87 | 0.27 | 0.67 | 0.54
Results from laboratories 1.11, 1.13 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 241b (Certified Value 1.05 ppm) | [32 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.8 | 1.107 | 1.589 | 0.72262 | 1.0912 | 1.1 | 1.1 | 1.28 | 1.05 | 1.1 | 1 | 1 | 1 | < 100 | 1 | 0.91 | 0.9 | 1.2 | 1.1 | 1.1 | 1.5 | 1 | 1.1 | 1 | [7 empty] | 1
| 1 | 2 | 0.7 | 1.106 | 1.023 | 0.71162 | 1.085 | 1.1 | 1.12 | 1.25 | 1.09 | 1.1 | 1 | 1 | 1 | < 100 | 1.1 | 0.96 | 0.9 | 1.1 | 1.1 | 1.1 | 1.5 | 0.9 | 1.1 | 1.1 | [7 empty] | 24
| 1 | 3 | 0.8 | 1.111 | 1.046 | 0.71307 | 1.0864 | 1.1 | 1.09 | 1.22 | 1.06 | 1.2 | 1 | 1 | 1 | < 100 | 1 | 1.05 | 0.9 | 1.2 | 1.1 | 1.1 | 1.5 | 1 | 1 | 1.1 | [7 empty] | 16
| 1 | 4 | 0.7 | 1.113 | 1.794 | 0.72145 | 1.0895 | 1.1 | 1.14 | 1.24 | 1.07 | 1.2 | 1 | 1 | 1.1 | < 100 | 1 | 0.97 | 0.8 | 1.1 | 1 | 1 | 1.4 | 1 | 1 | 1 | [7 empty] | 1.0542
| 1 | 5 | 0.8 | 1.111 | 1.125 | 0.72183 | 1.109 | 1.1 | 1.1 | 1.29 | 1.01 | 1.2 | 1 | 1 | 1 | < 100 | 1 | 0.93 | 0.9 | 1.1 | 1 | 1.1 | 1.4 | 1 | 1 | 1 | [7 empty] | 121
| 1 | 6 | 0.7 | 1.109 | 1.082 | 0.71045 | 1.0796 | 1 | 1.21 | 1.28 | 1.02 | 1.2 | 1 | 1 | 1 | < 100 | 1 | 1.04 | 0.9 | 1.1 | 1.1 | 1.1 | 1.4 | 1 | 1 | 1.1
| Mean |  | 0.75 | 1.1095 | 1.2765 | 0.71684 | 1.0901 | 1.0833 | 1.1267 | 1.26 | 1.05 | 1.1667 | 1 | 1 | 1.0167 |  | 1.0167 | 0.97667 | 0.88333 | 1.1333 | 1.0667 | 1.0833 | 1.45 | 0.98333 | 1.0333 | 1.05
| Median |  | 0.75 | 1.11 | 1.1035 | 0.71726 | 1.0879 | 1.1 | 1.11 | 1.265 | 1.055 | 1.2 | 1 | 1 | 1 |  | 1 | 0.965 | 0.9 | 1.1 | 1.1 | 1.1 | 1.45 | 1 | 1 | 1.05
| Std Dev. |  | 0.054772 | 0.0026646 | 0.32974 | 0.0056888 | 0.010113 | 0.040825 | 0.044572 | 0.027568 | 0.030332 | 0.05164 | 0 | 0 | 0.040825 |  | 0.040825 | 0.057155 | 0.040825 | 0.05164 | 0.05164 | 0.040825 | 0.054772 | 0.040825 | 0.05164 | 0.054772
| Rel.Std.Dev. |  | 7.303% | 0.2402% | 25.83% | 0.7936% | 0.9277% | 3.768% | 3.956% | 2.188% | 2.889% | 4.426% | 0% | 0% | 4.016% |  | 4.016% | 5.852% | 4.622% | 4.556% | 4.841% | 3.768% | 3.777% | 4.152% | 4.997% | 5.216%
| PDM3 |  | -28.86% | 5.242% | 21.08% | -32% | 3.404% | 2.76% | 6.871% | 19.52% | -0.4014% | 10.67% | -5.144% | -5.144% | -3.563% |  | -3.563% | -7.358% | -16.21% | 7.503% | 1.179% | 2.76% | 37.54% | -6.725% | -1.982% | -0.4014%
| Z-Score (Absolute) |  | 3.76 | 0.4 | 2.33 | 4.14 | 0.17 | 0.1 | 0.6 | 2.14 | 0.29 | 1.06 | NA | NA | 0.67 | 565.55 | 0.67 | 1.14 | 2.22 | 0.67 | 0.1 | 0.1 | 4.33 | 1.06 | 0.48 | 0.29
Results from laboratories 1.14 and 1.15 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ti in OREAS 241b (Certified Value 0.302 wt.%) | [32 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 3
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 3
| 1 | 1 | 0.33 | 0.2838 | 0.30333 | 0.41864 | 0.31969 | 0.38 | 0.3849 | 0.436 | 0.3697 | 0.3 | 0.255 | 0.3 | 0.314 | 0.297 | < 0.01 | 0.292 | 0.25 | 0.36463 | 0.4002 | 0.226 | 0.29 | 0.25021 | 0.30972 | 0.112 | 0.292 | 0.242 | 0.31 | 0.286 | 0.217 |  |  | 1
| 1 | 2 | 0.31 | 0.2852 | 0.31287 | 0.41459 | 0.32682 | 0.38 | 0.3921 | 0.425 | 0.3797 | 0.3 | 0.259 | 0.288 | 0.312 | 0.292 | < 0.01 | 0.287 | 0.25 | 0.42266 | 0.3721 | 0.273 | 0.29 | 0.2719 | 0.31366 | 0.1175 | 0.298 | 0.26 | 0.3 | 0.274 | 0.209 |  |  | 25
| 1 | 3 | 0.3 | 0.2754 | 0.33243 | 0.4093 | 0.32078 | 0.39 | 0.3879 | 0.417 | 0.377 | 0.31 | 0.259 | 0.295 | 0.301 | 0.29 | < 0.01 | 0.288 | 0.25 | 0.41681 | 0.3544 | 0.237 | 0.3 | 0.25338 | 0.31294 | 0.1155 | 0.295 | 0.263 | 0.3 | 0.273 | 0.22 |  |  | 16
| 1 | 4 | 0.3 | 0.2815 | 0.29924 | 0.41641 | 0.31638 | 0.38 | 0.3909 | 0.423 | 0.3708 | 0.31 | 0.259 | 0.292 | 0.31 | 0.298 | < 0.01 | 0.286 | 0.25 | 0.37151 | 0.3818 | 0.223 | 0.29 | 0.26152 | 0.29026 | 0.1045 | 0.29 | 0.268 | 0.3 | 0.28 | 0.206 |  |  | 0.30233
| 1 | 5 | 0.31 | 0.288 | 0.32898 | 0.41153 | 0.32698 | 0.37 | 0.3993 | 0.433 | 0.3608 | 0.3 | 0.262 | 0.293 | 0.309 | 0.294 | < 0.01 | 0.272 | 0.23 | 0.36034 | 0.3933 | 0.26 | 0.3 | 0.24262 | 0.29793 | 0.1075 | 0.297 | 0.275 | 0.29 | 0.285 | 0.216 |  |  | 122
| 1 | 6 | 0.29 | 0.2801 | 0.32137 | 0.42017 | 0.31755 | 0.39 | 0.4029 | 0.44 | 0.3657 | 0.3 | 0.263 | 0.295 | 0.315 | 0.285 | < 0.01 | 0.277 | 0.23 | 0.36074 | 0.3555 | 0.253 | 0.28 | 0.2324 | 0.30149 | 0.1165 | 0.294 | 0.217 | 0.28 | 0.279 | 0.224
| Mean |  | 0.30667 | 0.28233 | 0.31637 | 0.41511 | 0.32137 | 0.38167 | 0.393 | 0.429 | 0.37062 | 0.30333 | 0.2595 | 0.29383 | 0.31017 | 0.29267 |  | 0.28367 | 0.24333 | 0.38278 | 0.37622 | 0.24533 | 0.29167 | 0.25201 | 0.30433 | 0.11225 | 0.29433 | 0.25417 | 0.29667 | 0.2795 | 0.21533
| Median |  | 0.305 | 0.28265 | 0.31712 | 0.4155 | 0.32023 | 0.38 | 0.3915 | 0.429 | 0.37025 | 0.3 | 0.259 | 0.294 | 0.311 | 0.293 |  | 0.2865 | 0.25 | 0.36807 | 0.37695 | 0.245 | 0.29 | 0.25179 | 0.30561 | 0.11375 | 0.2945 | 0.2615 | 0.3 | 0.2795 | 0.2165
| Std Dev. |  | 0.013663 | 0.0043871 | 0.013549 | 0.0041617 | 0.0045581 | 0.0075277 | 0.0068489 | 0.0087407 | 0.0069953 | 0.005164 | 0.0028107 | 0.0039707 | 0.0050365 | 0.0048028 |  | 0.0075542 | 0.010328 | 0.028962 | 0.01908 | 0.019906 | 0.0075277 | 0.013883 | 0.0093512 | 0.0052702 | 0.0030111 | 0.021292 | 0.010328 | 0.0053944 | 0.0067429
| Rel.Std.Dev. |  | 4.455% | 1.554% | 4.283% | 1.003% | 1.418% | 1.972% | 1.743% | 2.037% | 1.887% | 1.702% | 1.083% | 1.351% | 1.624% | 1.641% |  | 2.663% | 4.244% | 7.566% | 5.072% | 8.114% | 2.581% | 5.509% | 3.073% | 4.695% | 1.023% | 8.377% | 3.481% | 1.93% | 3.131%
| PDM3 |  | 1.434% | -6.614% | 4.643% | 37.3% | 6.296% | 26.24% | 29.99% | 41.9% | 22.59% | 0.3319% | -14.17% | -2.81% | 2.592% | -3.196% |  | -6.173% | -19.51% | 26.61% | 24.44% | -18.85% | -3.527% | -16.65% | 0.6627% | -62.87% | -2.645% | -15.93% | -1.873% | -7.551% | -28.78%
| Z-Score (Absolute) |  | 0.24 | 0.23 | 0.43 | 2.34 | 0.52 | 1.69 | 1.91 | 2.61 | 1.48 | 0.17 | 0.67 | 0.01 | 0.31 | 0.03 | 5.6 | 0.21 | 0.99 | 1.71 | 1.59 | 0.95 | 0.05 | 0.82 | 0.19 | 3.52 | 0 | 0.78 | 0.05 | 0.29 | 1.53
| Analytical results for Tl in OREAS 241b (Certified Value 0.16 ppm) | [32 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.07 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.25 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.25g | 0.5g | 0.2g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.15g | 01g | 15g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.16 | 0.17 | 0.17971 | 0.15 | 0.16 | 0.17 | < 10 | 0.19 | 0.18 | 0.16 | 0.16 | 0.16 | < 5 | 0.16 | 0.16 | 0.17 | 5.49 | 0.15 | 0.17262 | 0.17 | 0.164 | 0.08 | 0.16 | 0.17 | 0.16 | [6 empty] | 1
| 1 | 2 | 0.17 | 0.18 | 0.16958 | 0.16 | 0.17 | 0.16 | < 10 | 0.19 | 0.18 | 0.16 | 0.16 | 0.16 | < 5 | 0.17 | 0.16 | 0.17 | 3.8 | 0.15 | 0.16364 | 0.17 | 0.233 | 0.08 | 0.15 | 0.16 | 0.16 | [6 empty] | 26
| 1 | 3 | 0.16 | 0.16 | 0.16894 | 0.17 | 0.17 | 0.16 | < 10 | 0.2 | 0.17 | 0.16 | 0.16 | 0.16 | < 5 | 0.17 | 0.16 | 0.17 | 5.32 | 0.15 | 0.16158 | 0.16 | 0.173 | 0.09 | 0.16 | 0.17 | 0.16 | [6 empty] | 16
| 1 | 4 | 0.17 | 0.17 | 0.1637 | 0.17 | 0.17 | 0.16 | < 10 | 0.18 | 0.16 | 0.16 | 0.16 | 0.16 | < 5 | 0.15 | 0.15 | 0.16 | 8.13 | 0.15 | 0.1672 | 0.16 | 0.183 | 0.08 | 0.16 | 0.16 | 0.15 | [6 empty] | 0.16316
| 1 | 5 | 0.17 | 0.17 | 0.16625 | 0.16 | 0.18 | 0.15 | < 10 | 0.18 | 0.18 | 0.16 | 0.15 | 0.15 | < 5 | 0.15 | 0.17 | 0.17 | 11.63 | 0.15 | 0.16573 | 0.17 | 0.175 | 0.08 | 0.16 | 0.16 | 0.15 | [6 empty] | 123
| 1 | 6 | 0.17 | 0.17 | 0.16282 | 0.15 | 0.18 | 0.15 | < 10 | 0.18 | 0.17 | 0.15 | 0.16 | 0.16 | < 5 | 0.16 | 0.16 | 0.17 | 5.68 | 0.15 | 0.17178 | 0.16 | 0.176 | 0.08 | 0.15 | 0.15 | 0.16
| Mean |  | 0.16667 | 0.17 | 0.1685 | 0.16 | 0.17167 | 0.15833 |  | 0.18667 | 0.17333 | 0.15833 | 0.15833 | 0.15833 |  | 0.16 | 0.16 | 0.16833 | 6.675 | 0.15 | 0.16709 | 0.165 | 0.184 | 0.081667 | 0.15667 | 0.16167 | 0.15667
| Median |  | 0.17 | 0.17 | 0.1676 | 0.16 | 0.17 | 0.16 |  | 0.185 | 0.175 | 0.16 | 0.16 | 0.16 |  | 0.16 | 0.16 | 0.17 | 5.585 | 0.15 | 0.16646 | 0.165 | 0.1755 | 0.08 | 0.16 | 0.16 | 0.16
| Std Dev. |  | 0.005164 | 0.0063246 | 0.0061224 | 0.0089443 | 0.0075277 | 0.0075277 |  | 0.008165 | 0.008165 | 0.0040825 | 0.0040825 | 0.0040825 |  | 0.0089443 | 0.0063246 | 0.0040825 | 2.7987 | 0 | 0.0043978 | 0.0054772 | 0.024771 | 0.0040825 | 0.005164 | 0.0075277 | 0.005164
| Rel.Std.Dev. |  | 3.098% | 3.72% | 3.633% | 5.59% | 4.385% | 4.754% |  | 4.374% | 4.711% | 2.578% | 2.578% | 2.578% |  | 5.59% | 3.953% | 2.425% | 41.93% | 0% | 2.632% | 3.32% | 13.46% | 4.999% | 3.296% | 4.656% | 3.296%
| PDM3 |  | 2.152% | 4.195% | 3.276% | -1.935% | 5.216% | -2.956% |  | 14.41% | 6.238% | -2.956% | -2.956% | -2.956% |  | -1.935% | -1.935% | 3.173% | 3991% | -8.064% | 2.412% | 1.13% | 12.78% | -49.95% | -3.978% | -0.913% | -3.978%
| Z-Score (Absolute) |  | 0.17 | 0.51 | 0.35 | 0.51 | 0.67 | 0.67 | 489.04 | 2.19 | 0.84 | 0.67 | 0.67 | 0.67 | 236.18 | 0.51 | 0.51 | 0.34 | 658.46 | 1.52 | 0.21 | 0 | 1.92 | 8.43 | 0.84 | 0.34 | 0.84
| Analytical results for Tm in OREAS 241b (Certified Value 0.17 ppm) | [32 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [26 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.07 | 1.27 | [26 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | [26 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 25g | [26 empty] | 2
| 1 | 1 | 0.174 | 0.156 | 0.1466 | 0.2 | 0.15 | [26 empty] | 1
| 1 | 2 | 0.181 | 0.166 | 0.14718 | 0.2 | 0.15 | [26 empty] | 27
| 1 | 3 | 0.178 | 0.155 | 0.14641 | 0.2 | 0.15 | [26 empty] | 16
| 1 | 4 | 0.18 | 0.16 | 0.15415 | 0.2 | 0.15 | [26 empty] | 0.16728
| 1 | 5 | 0.18 | 0.17 | 0.14609 | 0.2 | 0.15 | [26 empty] | 124
| 1 | 6 | 0.173 | 0.165 | 0.14736 | 0.2 | 0.15
| Mean |  | 0.17767 | 0.162 | 0.14796 | 0.2 | 0.15
| Median |  | 0.179 | 0.1625 | 0.14689 | 0.2 | 0.15
| Std Dev. |  | 0.0033862 | 0.0059666 | 0.0030693 | 3.0405e-17 | 0
| Rel.Std.Dev. |  | 1.906% | 3.683% | 2.074% | 1.52e-14% | 0%
| PDM3 |  | 6.21% | -3.156% | -11.55% | 19.56% | -10.33%
| Z-Score (Absolute) |  | 0.75 | 0 | 0.67 | 1.83 | 0.58
| Analytical results for U in OREAS 241b (Certified Value 0.26 ppm) | [32 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [5 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [5 empty] | 3
| 1 | 1 | 0.28 | 0.271 | 0.369 | 0.18322 | 0.29335 | 0.3 | 0.29 | 0.33 | 0.29 | 0.3 | 0.23 | 0.26 | 0.25 | 1623.415 | 0.26 | < 10 | 0.25 | 0.2 | 0.28 | 0.2 | 0.26 | 0.24 | 0.22 | 0.26 | 0.27 | 0.24 | [5 empty] | 1
| 1 | 2 | 0.27 | 0.271 | 0.275 | 0.18916 | 0.29546 | 0.3 | 0.28 | 0.33 | 0.29 | 0.3 | 0.22 | 0.25 | 0.24 | 1649.43 | 0.26 | < 10 | 0.25 | 0.2 | 0.27 | 0.2 | 0.26 | 0.249 | 0.23 | 0.25 | 0.26 | 0.24 | [5 empty] | 28
| 1 | 3 | 0.26 | 0.264 | 0.282 | 0.18131 | 0.34646 | 0.3 | 0.29 | 0.34 | 0.29 | 0.3 | 0.24 | 0.25 | 0.24 | 1534.37 | 0.26 | < 10 | 0.26 | 0.2 | 0.26 | 0.2 | 0.25 | 0.248 | 0.23 | 0.26 | 0.25 | 0.24 | [5 empty] | 16
| 1 | 4 | 0.27 | 0.271 | 0.464 | 0.17678 | 0.28967 | 0.3 | 0.3 | 0.34 | 0.29 | 0.3 | 0.23 | 0.25 | 0.26 | 1595.764 | 0.26 | < 10 | 0.26 | 0.2 | 0.27 | 0.2 | 0.25 | 0.246 | 0.23 | 0.25 | 0.27 | 0.22 | [5 empty] | 0.25982
| 1 | 5 | 0.27 | 0.268 | 0.308 | 0.18894 | 0.29291 | 0.3 | 0.29 | 0.34 | 0.28 | 0.3 | 0.23 | 0.25 | 0.23 | 1467.665 | 0.25 | < 10 | 0.24 | 0.2 | 0.28 | 0.2 | 0.26 | 0.234 | 0.23 | 0.26 | 0.26 | 0.24 | [5 empty] | 125
| 1 | 6 | 0.27 | 0.274 | 0.297 | 0.18658 | 0.27888 | 0.3 | 0.28 | 0.36 | 0.28 | 0.3 | 0.23 | 0.25 | 0.24 | 1614.642 | 0.25 | < 10 | 0.25 | 0.2 | 0.26 | 0.2 | 0.25 | 0.241 | 0.22 | 0.25 | 0.26 | 0.24
| Mean |  | 0.27 | 0.26983 | 0.3325 | 0.18433 | 0.29946 | 0.3 | 0.28833 | 0.34 | 0.28667 | 0.3 | 0.23 | 0.25167 | 0.24333 | 1580.881 | 0.25667 |  | 0.25167 | 0.2 | 0.27 | 0.2 | 0.255 | 0.243 | 0.22667 | 0.255 | 0.26167 | 0.23667
| Median |  | 0.27 | 0.271 | 0.3025 | 0.1849 | 0.29313 | 0.3 | 0.29 | 0.34 | 0.29 | 0.3 | 0.23 | 0.25 | 0.24 | 1605.203 | 0.26 |  | 0.25 | 0.2 | 0.27 | 0.2 | 0.255 | 0.2435 | 0.23 | 0.255 | 0.26 | 0.24
| Std Dev. |  | 0.0063246 | 0.0034303 | 0.072583 | 0.0048366 | 0.023769 | 0 | 0.0075277 | 0.010954 | 0.005164 | 0 | 0.0063246 | 0.0040825 | 0.010328 | 67.6026 | 0.005164 |  | 0.0075277 | 3.0405e-17 | 0.0089443 | 3.0405e-17 | 0.0054772 | 0.0057271 | 0.005164 | 0.0054772 | 0.0075277 | 0.008165
| Rel.Std.Dev. |  | 2.342% | 1.271% | 21.83% | 2.624% | 7.938% | 0% | 2.611% | 3.222% | 1.801% | 0% | 2.75% | 1.622% | 4.244% | 4.276% | 2.012% |  | 2.991% | 1.52e-14% | 3.313% | 1.52e-14% | 2.148% | 2.357% | 2.278% | 2.148% | 2.877% | 3.45%
| PDM3 |  | 3.919% | 3.855% | 27.97% | -29.05% | 15.26% | 15.47% | 10.98% | 30.86% | 10.33% | 15.47% | -11.48% | -3.137% | -6.345% | 608358% | -1.213% |  | -3.137% | -23.02% | 3.919% | -23.02% | -1.854% | -6.473% | -12.76% | -1.854% | 0.7116% | -8.911%
| Z-Score (Absolute) |  | 0.38 | 0.37 | 2.56 | 2.61 | 1.41 | NA | 1.02 | 2.82 | 0.96 | NA | 1.02 | 0.26 | 0.55 | 55129 | 0.09 | 165.35 | 0.26 | NA | 0.38 | NA | 0.15 | 0.56 | 1.13 | 0.15 | 0.09 | 0.78
Results from laboratories 1.07, 1.13, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for V in OREAS 241b (Certified Value 131 ppm) | [32 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g |  |  | 0
| 1 | 1 | 139 | 137 | 110.68 | 129.534 | 139.3331 | 138 | 151 | 174 | 161 | 132 | 129 | 129 | 129 | 125 | < 10 | 128 | 116 | 157.2266 | 146 | 114 | 128 | 119.27 | 144.9736 | 90 | 129 | 110 | 126 | 139 | 102 |  |  | 1
| 1 | 2 | 136 | 137.9 | 113.64 | 133.2485 | 143.638 | 146 | 156 | 170 | 166 | 134 | 127 | 123 | 129 | 124 | < 10 | 127 | 117 | 166.2462 | 141 | 122 | 128 | 127.53 | 142.232 | 90 | 129 | 117 | 125 | 133 | 100 |  |  | 29
| 1 | 3 | 135 | 136.3 | 117.92 | 129.9247 | 138.8739 | 149 | 165 | 170 | 161 | 134 | 130 | 125 | 125 | 123 | < 10 | 127 | 117 | 162.7262 | 141 | 119 | 127 | 119.99 | 140.1988 | 95 | 129 | 117 | 125 | 132 | 104 |  |  | 16
| 1 | 4 | 136 | 138 | 109.62 | 132.9599 | 138.6854 | 140 | 158 | 170 | 161 | 136 | 128 | 124 | 129 | 125 | < 10 | 127 | 117 | 155.8362 | 141 | 110 | 128 | 122.95 | 128.7172 | 90 | 127 | 117 | 127 | 133 | 99.1 |  |  | 131.1176
| 1 | 5 | 136 | 136.9 | 116.56 | 133.1798 | 141.8172 | 139 | 158 | 174 | 156 | 135 | 129 | 125 | 128 | 125 | < 10 | 121 | 112 | 151.1406 | 146 | 122 | 128 | 116.89 | 128.8816 | 90 | 129 | 126 | 122 | 134 | 102 |  |  | 126
| 1 | 6 | 131 | 135.8 | 113.67 | 133.3554 | 137.6069 | 136 | 164 | 178 | 156 | 135 | 132 | 126 | 129 | 123 | < 10 | 124 | 111 | 152.562 | 141 | 118 | 126 | 113.28 | 140.65 | 95 | 128 | 106 | 125 | 133 | 102
| Mean |  | 135.5 | 136.9833 | 113.6817 | 132.0337 | 139.9924 | 141.3333 | 158.6667 | 172.6667 | 160.1667 | 134.3333 | 129.1667 | 125.3333 | 128.1667 | 124.1667 |  | 125.6667 | 115 | 157.6229 | 142.6667 | 117.5 | 127.5 | 119.985 | 137.6089 | 91.6667 | 128.5 | 115.5 | 125 | 134 | 101.5167
| Median |  | 136 | 136.95 | 113.655 | 133.0699 | 139.1035 | 139.5 | 158 | 172 | 161 | 134.5 | 129 | 125 | 129 | 124.5 |  | 127 | 116.5 | 156.5314 | 141 | 118.5 | 128 | 119.63 | 140.4244 | 90 | 129 | 117 | 125 | 133 | 102
| Std Dev. |  | 2.5884 | 0.86583 | 3.2174 | 1.7939 | 2.2674 | 5.0465 | 5.2026 | 3.266 | 3.7639 | 1.3663 | 1.7224 | 2.0656 | 1.6021 | 0.98319 |  | 2.6583 | 2.7568 | 5.8547 | 2.582 | 4.7223 | 0.83666 | 4.9146 | 7.0256 | 2.582 | 0.83666 | 6.892 | 1.6733 | 2.5298 | 1.7325
| Rel.Std.Dev. |  | 1.91% | 0.6321% | 2.83% | 1.359% | 1.62% | 3.571% | 3.279% | 1.891% | 2.35% | 1.017% | 1.333% | 1.648% | 1.25% | 0.7918% |  | 2.115% | 2.397% | 3.714% | 1.81% | 4.019% | 0.6562% | 4.096% | 5.105% | 2.817% | 0.6511% | 5.967% | 1.339% | 1.888% | 1.707%
| PDM3 |  | 3.342% | 4.474% | -13.3% | 0.6987% | 6.769% | 7.791% | 21.01% | 31.69% | 22.15% | 2.453% | -1.488% | -4.412% | -2.251% | -5.301% |  | -4.157% | -12.29% | 20.21% | 8.808% | -10.39% | -2.759% | -8.491% | 4.951% | -30.09% | -1.996% | -11.91% | -4.666% | 2.198% | -22.58%
| Z-Score (Absolute) |  | 0.52 | 0.63 | 1.1 | 0.26 | 0.85 | 0.95 | 2.23 | 3.27 | 2.34 | 0.43 | 0.05 | 0.23 | 0.02 | 0.32 | 9.14 | 0.21 | 1 | 2.16 | 1.05 | 0.81 | 0.07 | 0.63 | 0.67 | 2.73 | 0 | 0.96 | 0.26 | 0.41 | 2
| Analytical results for W in OREAS 241b (Certified Value 27.1 ppm) | [32 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.15g | 01g | 15g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [6 empty] | 1
| 1 | 1 | 31.6 | 24.54 | 33.39 | 33.587 | 22.4 | 33 | 19.9 | 19.91 | 23 | 28.7 | 30.4 | 30.7 | 28.1 | < 100 | 28.3 | 21 | 26.4 | 20.7 | 29 | 21 | 29.2 | 32.2 | 31.1 | 23.37 | 32.9 | [6 empty] | 1
| 1 | 2 | 31.2 | 24.53 | 32.64 | 33.0856 | 23.8 | 34 | 19.2 | 20.01 | 23 | 28.6 | 29.2 | 31 | 27.5 | < 100 | 27.3 | 20 | 27.1 | 21.6 | 29 | 22 | 28.8 | 32.4 | 29.5 | 23.48 | 33.4 | [6 empty] | 30
| 1 | 3 | 30.8 | 24.26 | 34.05 | 30.1571 | 23.9 | 33 | 20.3 | 18.18 | 23 | 28.7 | 30.1 | 30 | 27.6 | < 100 | 27.3 | 20 | 30.5 | 21 | 29 | 21 | 29.4 | 32.9 | 30.2 | 23.05 | 33.5 | [6 empty] | 16
| 1 | 4 | 31.9 | 25.13 | 32.55 | 33.1372 | 21.6 | 31 | 19.1 | 18.7 | 23 | 28.4 | 29 | 29.6 | 28 | < 100 | 27.5 | 21 | 28.8 | 19.3 | 29 | 22 | 29.2 | 32 | 30 | 23.5 | 32 | [6 empty] | 27.1493
| 1 | 5 | 31.7 | 24.86 | 34.75 | 32.8683 | 23.2 | 30 | 19.1 | 19.92 | 23 | 28.9 | 30 | 29.1 | 27.1 | < 100 | 27.3 | 21 | 27.7 | 20.2 | 29 | 22 | 30.1 | 32.2 | 30.8 | 24.03 | 33 | [6 empty] | 127
| 1 | 6 | 32.5 | 24.41 | 33.91 | 32.9607 | 22.7 | 30 | 19.2 | 19.23 | 23 | 29.2 | 29.4 | 29.6 | 26.9 | < 100 | 27.4 | 21 | 29.9 | 20.7 | 29 | 22 | 29.9 | 33.4 | 30.3 | 23.17 | 32.8
| Mean |  | 31.6167 | 24.6217 | 33.5483 | 32.6327 | 22.9333 | 31.8333 | 19.4667 | 19.325 | 23 | 28.75 | 29.6833 | 30 | 27.5333 |  | 27.5167 | 20.6667 | 28.4 | 20.5833 | 29 | 21.6667 | 29.4333 | 32.5167 | 30.3167 | 23.4333 | 32.9333
| Median |  | 31.65 | 24.535 | 33.65 | 33.0232 | 22.95 | 32 | 19.2 | 19.57 | 23 | 28.7 | 29.7 | 29.8 | 27.55 |  | 27.35 | 21 | 28.25 | 20.7 | 29 | 22 | 29.3 | 32.3 | 30.25 | 23.425 | 32.95
| Std Dev. |  | 0.58452 | 0.31808 | 0.85712 | 1.2379 | 0.88015 | 1.7224 | 0.50859 | 0.75844 | 0 | 0.27386 | 0.56006 | 0.72388 | 0.4761 |  | 0.392 | 0.5164 | 1.6125 | 0.77825 | 0 | 0.5164 | 0.48442 | 0.53072 | 0.57067 | 0.34145 | 0.53541
| Rel.Std.Dev. |  | 1.849% | 1.292% | 2.555% | 3.793% | 3.838% | 5.411% | 2.613% | 3.925% | 0% | 0.9526% | 1.887% | 2.413% | 1.729% |  | 1.425% | 2.499% | 5.678% | 3.781% | 0% | 2.383% | 1.646% | 1.632% | 1.882% | 1.457% | 1.626%
| PDM3 |  | 16.45% | -9.31% | 23.57% | 20.2% | -15.53% | 17.25% | -28.3% | -28.82% | -15.28% | 5.896% | 9.334% | 10.5% | 1.415% |  | 1.353% | -23.88% | 4.607% | -24.18% | 6.817% | -20.19% | 8.413% | 19.77% | 11.67% | -13.69% | 21.3%
| Z-Score (Absolute) |  | 0.5 | 0.72 | 0.83 | 0.67 | 1.01 | 0.54 | 1.61 | 1.64 | 1 | 0 | 0.16 | 0.22 | 0.21 | 3.69 | 0.21 | 1.4 | 0.06 | 1.42 | 0.04 | 1.23 | 0.12 | 0.65 | 0.27 | 0.92 | 0.73
| Analytical results for Y in OREAS 241b (Certified Value 10.9 ppm) | [32 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 2
| 1 | 1 | 11.28 | 11.88 | 10.39 | 10.5574 | 11.2155 | 11.5 | 11.1 | 14.6 | 11.5 | 10 | 11.2 | 10.75 | 10.75 | 10.45 | 11.35 | 10 | 12.05 | 11 | 9.3 | 12 | 9.55 | 10.8 | 11.77 | 9.7 | [7 empty] | 1
| 1 | 2 | 11.14 | 11.82 | 10.63 | 10.4945 | 11.5548 | 12.2 | 11 | 14.3 | 11.76 | 10 | 11.1 | 10.85 | 10.55 | 10.5 | 11.45 | 10 | 11.77 | 11 | 9.3 | 12 | 9.9 | 10.2 | 11.24 | 9.65 | [7 empty] | 31
| 1 | 3 | 11.12 | 11.73 | 10.86 | 10.5497 | 11.1779 | 12.1 | 11 | 14.5 | 11.14 | 10 | 11.4 | 10.95 | 10.8 | 10.45 | 11.25 | 10 | 11.71 | 11.1 | 9.3 | 12.05 | 10 | 10.5 | 10.84 | 9.9 | [7 empty] | 16
| 1 | 4 | 11.31 | 11.93 | 10.3 | 10.5749 | 11.2715 | 11.7 | 11.2 | 14.2 | 11.17 | 10 | 11.2 | 10.9 | 11 | 10.65 | 11.2 | 10 | 11.93 | 10.9 | 8.9 | 11.55 | 10.2 | 10.6 | 11.04 | 9.55 | [7 empty] | 10.9312
| 1 | 5 | 11.26 | 12.01 | 10.76 | 10.7172 | 11.3871 | 12 | 11.2 | 14.6 | 10.95 | 10 | 11.3 | 10.85 | 10.55 | 10.4 | 10.85 | 9 | 11.98 | 11.1 | 9.3 | 12.25 | 10.1 | 10.7 | 11 | 9.81 | [7 empty] | 128
| 1 | 6 | 10.93 | 11.9 | 10.53 | 10.7695 | 11.0237 | 11.2 | 11.4 | 14.8 | 10.97 | 10 | 11.4 | 10.9 | 10.8 | 10.55 | 11 | 9 | 11.76 | 11 | 9.2 | 11.9 | 9.38 | 10.2 | 11.02 | 9.69
| Mean |  | 11.1733 | 11.8783 | 10.5783 | 10.6106 | 11.2718 | 11.7833 | 11.15 | 14.5 | 11.2483 | 10 | 11.2667 | 10.8667 | 10.7417 | 10.5 | 11.1833 | 9.6667 | 11.8667 | 11.0167 | 9.2167 | 11.9583 | 9.855 | 10.5 | 11.1517 | 9.7167
| Median |  | 11.2 | 11.89 | 10.58 | 10.5662 | 11.2435 | 11.85 | 11.15 | 14.55 | 11.155 | 10 | 11.25 | 10.875 | 10.775 | 10.475 | 11.225 | 10 | 11.85 | 11 | 9.3 | 12 | 9.95 | 10.55 | 11.03 | 9.695
| Std Dev. |  | 0.14194 | 0.095795 | 0.21461 | 0.10763 | 0.18269 | 0.38687 | 0.15166 | 0.21909 | 0.31921 | 0 | 0.12111 | 0.068313 | 0.17151 | 0.089443 | 0.22286 | 0.5164 | 0.13837 | 0.075277 | 0.16021 | 0.23112 | 0.32272 | 0.25298 | 0.32866 | 0.12291
| Rel.Std.Dev. |  | 1.27% | 0.8065% | 2.029% | 1.014% | 1.621% | 3.283% | 1.36% | 1.511% | 2.838% | 0% | 1.075% | 0.6286% | 1.597% | 0.8518% | 1.993% | 5.342% | 1.166% | 0.6833% | 1.738% | 1.933% | 3.275% | 2.409% | 2.947% | 1.265%
| PDM3 |  | 2.215% | 8.664% | -3.228% | -2.933% | 3.115% | 7.795% | 2.002% | 32.65% | 2.901% | -8.519% | 3.069% | -0.5904% | -1.734% | -3.945% | 2.306% | -11.57% | 8.558% | 0.7818% | -15.68% | 9.396% | -9.845% | -3.945% | 2.017% | -11.11%
| Z-Score (Absolute) |  | 0.03 | 0.88 | 0.69 | 0.65 | 0.15 | 0.77 | 0 | 4.06 | 0.12 | NA | 0.14 | 0.34 | 0.5 | 0.79 | 0.04 | NA | 0.87 | 0.16 | 2.34 | 0.98 | 1.57 | 0.79 | 0 | 1.74
Results from laboratories 1.12 and 1.19 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 241b (Certified Value 1.03 ppm) | [32 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [23 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.07 | 1.08 | 1.27 | 1.32 | [23 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | [23 empty] | 2
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.25g | [23 empty] | 2
| 1 | 1 | 1.1 | 1.101 | 0.871 | 0.92818 | 1.2 | 1.22 | 0.9 | 1.1 | [23 empty] | 1
| 1 | 2 | 1 | 1.079 | 0.922 | 0.92274 | 1.2 | 1.24 | 0.85 | 1 | [23 empty] | 32
| 1 | 3 | 1.1 | 1.092 | 0.865 | 0.91008 | 1.2 | 1.3 | 0.85 | 1 | [23 empty] | 16
| 1 | 4 | 1.1 | 1.12 | 0.898 | 0.93538 | 1.2 | 1.24 | 0.85 | 0.9 | [23 empty] | 1.0321
| 1 | 5 | 1.1 | 1.105 | 0.905 | 0.90752 | 1.2 | 1.24 | 0.85 | 1 | [23 empty] | 129
| 1 | 6 | 1 | 1.135 | 0.9 | 0.9253 | 1.1 | 1.28 | 0.85 | 0.9
| Mean |  | 1.0667 | 1.1053 | 0.8935 | 0.92153 | 1.1833 | 1.2533 | 0.85833 | 0.98333
| Median |  | 1.1 | 1.103 | 0.899 | 0.92402 | 1.2 | 1.24 | 0.85 | 1
| Std Dev. |  | 0.05164 | 0.019927 | 0.021566 | 0.010761 | 0.040825 | 0.030111 | 0.020412 | 0.075277
| Rel.Std.Dev. |  | 4.841% | 1.803% | 2.414% | 1.168% | 3.45% | 2.402% | 2.378% | 7.655%
| PDM3 |  | 3.346% | 7.093% | -13.43% | -10.72% | 14.65% | 21.43% | -16.84% | -4.728%
| Z-Score (Absolute) |  | 0.24 | 0.46 | 0.75 | 0.59 | 0.91 | 1.31 | 0.96 | 0.24
| Analytical results for Zn in OREAS 241b (Certified Value 149 ppm) | [32 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.29 | 1.30 | 1.31 | 1.32 | 1.33 | 1.34 | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | 0
|  |  | 0.25g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.2g | 0.5g | 0.15g | 0.25g | 01g | 15g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | 0
| 1 | 1 | 153 | 147 | 139.43 | 141.41 | 151.6665 | 146 | 142 | 155 | 150 | 138 | 159 | 160 | 148 | 149 | < 100 | 163 | 136 | 142.7803 | 151 | 146.5 | 149 | 156.24 | 156.05 | 145 | 164.747 | 156 | 128.675 | 146 | 159 | 146 | 152 | 1
| 1 | 2 | 152 | 146.8 | 137.04 | 143.1 | 157.1891 | 155 | 150 | 152 | 155 | 139 | 155 | 154 | 148 | 151 | < 100 | 162 | 137 | 138.9221 | 157 | 149.7 | 140 | 154.04 | 155.61 | 145 | 166.588 | 155 | 153.883 | 151 | 152 | 143 | 152 | 33
| 1 | 3 | 152 | 147.2 | 140.16 | 140.85 | 152.332 | 158 | 148 | 154 | 149 | 138 | 161 | 155 | 148 | 150 | < 100 | 162 | 136 | 141.367 | 154 | 149.7 | 147 | 145.87 | 152.94 | 150 | 164.2618 | 157 | 138.804 | 150 | 157 | 144 | 156 | 16
| 1 | 4 | 154 | 146.4 | 132.59 | 143.45 | 154.6375 | 147 | 148 | 151 | 148 | 141 | 158 | 154 | 147 | 150 | < 100 | 160 | 138 | 144.8053 | 150 | 140.1 | 151 | 149.15 | 149.02 | 140 | 166.3164 | 154 | 139.684 | 146 | 159 | 143 | 150 | 149.3803
| 1 | 5 | 157 | 145.1 | 135.16 | 146.26 | 154.5957 | 151 | 142 | 156 | 148 | 140 | 158 | 155 | 150 | 152 | < 100 | 160 | 136 | 142.0733 | 152 | 150.5 | 148 | 147.51 | 147.72 | 145 | 162.7728 | 156 | 151.72 | 148 | 161 | 143 | 155 | 130
| 1 | 6 | 156 | 145.3 | 130.94 | 145.44 | 150.3476 | 147 | 148 | 157 | 145 | 140 | 161 | 154 | 149 | 151 | < 100 | 164 | 136 | 137.117 | 158 | 143.7 | 140 | 154.26 | 154.73 | 150 | 158.8819 | 156 | 133.916 | 145 | 154 | 141 | 151
| Mean |  | 154 | 146.3 | 135.8867 | 143.4183 | 153.4614 | 150.6667 | 146.3333 | 154.1667 | 149.1667 | 139.3333 | 158.6667 | 155.3333 | 148.3333 | 150.5 |  | 161.8333 | 136.5 | 141.1775 | 153.6667 | 146.7 | 145.8333 | 151.1783 | 152.6783 | 145.8333 | 163.928 | 155.6667 | 141.1137 | 147.6667 | 157 | 143.3333 | 152.6667
| Median |  | 153.5 | 146.6 | 136.1 | 143.275 | 153.4639 | 149 | 148 | 154.5 | 148.5 | 139.5 | 158.5 | 154.5 | 148 | 150.5 |  | 162 | 136 | 141.7201 | 153 | 148.1 | 147.5 | 151.595 | 153.835 | 145 | 164.5044 | 156 | 139.244 | 147 | 158 | 143 | 152
| Std Dev. |  | 2.0976 | 0.89443 | 3.6882 | 2.1398 | 2.4803 | 4.9261 | 3.4448 | 2.3166 | 3.3116 | 1.2111 | 2.2509 | 2.3381 | 1.0328 | 1.0488 |  | 1.6021 | 0.83666 | 2.762 | 3.266 | 4.1241 | 4.7081 | 4.2203 | 3.5273 | 3.7639 | 2.8411 | 1.0328 | 9.894 | 2.4221 | 3.4059 | 1.633 | 2.3381
| Rel.Std.Dev. |  | 1.362% | 0.6114% | 2.714% | 1.492% | 1.616% | 3.27% | 2.354% | 1.503% | 2.22% | 0.8692% | 1.419% | 1.505% | 0.6963% | 0.6969% |  | 0.99% | 0.6129% | 1.956% | 2.125% | 2.811% | 3.228% | 2.792% | 2.31% | 2.581% | 1.733% | 0.6635% | 7.011% | 1.64% | 2.169% | 1.139% | 1.532%
| PDM3 |  | 3.093% | -2.062% | -9.033% | -3.991% | 2.732% | 0.8611% | -2.04% | 3.204% | -0.143% | -6.726% | 6.217% | 3.985% | -0.7009% | 0.7496% |  | 8.336% | -8.622% | -5.491% | 2.869% | -1.794% | -2.374% | 1.204% | 2.208% | -2.374% | 9.739% | 4.208% | -5.534% | -1.147% | 5.101% | -4.048% | 2.2%
| Z-Score (Absolute) |  | 0.67 | 0.4 | 1.85 | 0.8 | 0.6 | 0.21 | 0.4 | 0.7 | 0 | 1.37 | 1.33 | 0.86 | 0.12 | 0.19 | 13.83 | 1.77 | 1.77 | 1.11 | 0.63 | 0.34 | 0.47 | 0.28 | 0.49 | 0.47 | 2.06 | 0.91 | 1.12 | 0.21 | 1.09 | 0.81 | 0.49
| Analytical results for Zr in OREAS 241b (Certified Value 16 ppm) | [32 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.07 | 1.08 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.19 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.31 | 1.32 | 1.33 | 1.34 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [7 empty] | 1
|  |  | 0.25g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 01g | 0.5g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.15g | 01g | 0.25g | 0.5g | 25g | 0.5g | 0.5g | 0.25g | 0.5g | 0.5g | [7 empty] | 1
| 1 | 1 | 15.3 | 16.71 | 20.55 | 16.9571 | 14.9 | 16.8 | 21.4 | 21.4 | 14 | 13.7 | 15.3 | 13 | 15.6 | 17 | 14 | 15 | 15.3 | 20.43 | 9 | 18.3 | 13 | 15.7 | 16.5 | 12.5 | [7 empty] | 1
| 1 | 2 | 15.1 | 17.17 | 20.38 | 17.3237 | 14.9 | 14.8 | 20.5 | 21.6 | 14 | 13.9 | 15.5 | 12.5 | 15.6 | 16.8 | 15 | 16.1 | 15 | 19.584 | 9 | 17.6 | 13.6 | 14.9 | 16.2 | 12 | [7 empty] | 34
| 1 | 3 | 14.1 | 17.11 | 21.52 | 16.7106 | 16.3 | 15.3 | 21.6 | 20.3 | 14 | 14.1 | 15.4 | 12.8 | 15 | 16.6 | 15 | 15.6 | 15.5 | 19.53 | 9 | 17.7 | 13.9 | 15.6 | 15.7 | 12.8 | [7 empty] | 16
| 1 | 4 | 14.3 | 16.68 | 20.15 | 16.8386 | 16.1 | 23.4 | 20.9 | 20.5 | 15 | 13.8 | 15.5 | 12.9 | 15.8 | 16.5 | 15 | 13.9 | 15.3 | 17.631 | 8 | 17.1 | 14.4 | 16 | 16.3 | 11.9 | [7 empty] | 15.9799
| 1 | 5 | 14.6 | 17.28 | 21.43 | 17.4702 | 15 | 14.7 | 21.2 | 20.6 | 14 | 13.7 | 15.8 | 12.5 | 15.2 | 15.9 | 14 | 13.9 | 15.3 | 18.498 | 9 | 18.3 | 14.7 | 15.4 | 16.4 | 12.6 | [7 empty] | 131
| 1 | 6 | 13.8 | 16.81 | 20.99 | 16.6892 | 14.6 | 15.3 | 21.7 | 20.5 | 14 | 14.3 | 15.6 | 12.9 | 15.1 | 15.9 | 13 | 14 | 15.1 | 18.808 | 10 | 17.9 | 12.4 | 14.9 | 16.4 | 12.6
| Mean |  | 14.5333 | 16.96 | 20.8367 | 16.9982 | 15.3 | 16.7167 | 21.2167 | 20.8167 | 14.1667 | 13.9167 | 15.5167 | 12.7667 | 15.3833 | 16.45 | 14.3333 | 14.75 | 15.25 | 19.0802 | 9 | 17.8167 | 13.6667 | 15.4167 | 16.25 | 12.4
| Median |  | 14.45 | 16.96 | 20.77 | 16.8978 | 14.95 | 15.3 | 21.3 | 20.55 | 14 | 13.85 | 15.5 | 12.85 | 15.4 | 16.55 | 14.5 | 14.5 | 15.3 | 19.169 | 9 | 17.8 | 13.75 | 15.5 | 16.35 | 12.55
| Std Dev. |  | 0.58195 | 0.25784 | 0.56652 | 0.32682 | 0.71274 | 3.3594 | 0.4535 | 0.54191 | 0.40825 | 0.24014 | 0.17224 | 0.21602 | 0.32506 | 0.45935 | 0.8165 | 0.96073 | 0.17607 | 0.97891 | 0.63246 | 0.45789 | 0.86178 | 0.4446 | 0.2881 | 0.36332
| Rel.Std.Dev. |  | 4.004% | 1.52% | 2.719% | 1.923% | 4.658% | 20.1% | 2.137% | 2.603% | 2.882% | 1.726% | 1.11% | 1.692% | 2.113% | 2.792% | 5.696% | 6.513% | 1.155% | 5.131% | 7.027% | 2.57% | 6.306% | 2.884% | 1.773% | 2.93%
| PDM3 |  | -9.053% | 6.133% | 30.39% | 6.372% | -4.255% | 4.61% | 32.77% | 30.27% | -11.35% | -12.91% | -2.899% | -20.11% | -3.733% | 2.942% | -10.3% | -7.697% | -4.568% | 19.4% | -43.68% | 11.49% | -14.48% | -3.525% | 1.69% | -22.4%
| Z-Score (Absolute) |  | 0.46 | 0.67 | 2.49 | 0.69 | 0.1 | 0.56 | 2.66 | 2.48 | NA | 0.75 | 0 | 1.28 | 0.06 | 0.44 | NA | 0.36 | 0.12 | 1.66 | NA | 1.07 | 0.86 | 0.05 | 0.34 | 1.46
Results from laboratories 1.12, 1.19 and 1.27 were removed due to their 1 ppm reading resolution.
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C146: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C164: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C183: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C220: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C238: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C440: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C458: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C476: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C494: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C512: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C550: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C568: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C623: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C839: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C858: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C877: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C895: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C913: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C931: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C949: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C967: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C985: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1042: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1151: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1170: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1188: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1206: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 241b (Indicative Value 12.39 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 12.39 | 1
| 1 | 2 | 12.39 | 9
| Mean |  | 12.39 | 16
| Median |  | 12.39 | 12.39
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 241b (Indicative Value 8.36 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 8.36 | 1
| 1 | 2 | 8.36 | 10
| Mean |  | 8.36 | 16
| Median |  | 8.36 | 8.36
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 241b (Indicative Value 11.16 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 11.15 | 1
| 1 | 2 | 11.16 | 11
| Mean |  | 11.155 | 16
| Median |  | 11.155 | 11.155
| Std Dev. |  | 0.0070711 | 17
| Rel.Std.Dev. |  | 0.06339%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 241b (Indicative Value 0.735 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.733 | 1
| 1 | 2 | 0.736 | 12
| Mean |  | 0.7345 | 16
| Median |  | 0.7345 | 0.7345
| Std Dev. |  | 0.0021213 | 18
| Rel.Std.Dev. |  | 0.2888%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 241b (Indicative Value 5.62 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.62 | 1
| 1 | 2 | 5.61 | 9
| Mean |  | 5.615 | 16
| Median |  | 5.615 | 5.615
| Std Dev. |  | 0.0070711 | 15
| Rel.Std.Dev. |  | 0.1259%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 241b (Indicative Value 0.162 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.162 | 1
| 1 | 2 | 0.161 | 10
| Mean |  | 0.1615 | 16
| Median |  | 0.1615 | 0.1615
| Std Dev. |  | 0.00070711 | 16
| Rel.Std.Dev. |  | 0.4378%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 241b (Indicative Value 2.59 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.6 | 1
| 1 | 2 | 2.58 | 11
| Mean |  | 2.59 | 16
| Median |  | 2.59 | 2.59
| Std Dev. |  | 0.014142 | 17
| Rel.Std.Dev. |  | 0.546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 241b (Indicative Value 0.097 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.097 | 1
| 1 | 2 | 0.096 | 12
| Mean |  | 0.0965 | 16
| Median |  | 0.0965 | 0.0965
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.7328%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 241b (Indicative Value 0.412 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.4125 | 1
| 1 | 2 | 0.4125 | 16
| Mean |  | 0.4125 | 16
| Median |  | 0.4125 | 0.41246
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.008849%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 241b (Indicative Value 54.4 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 54.43 | 1
| 1 | 2 | 54.36 | 10
| Mean |  | 54.395 | 16
| Median |  | 54.395 | 54.395
| Std Dev. |  | 0.049497 | 16
| Rel.Std.Dev. |  | 0.091%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 241b (Indicative Value 1.04 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.037 | 1
| 1 | 2 | 1.043 | 11
| Mean |  | 1.04 | 16
| Median |  | 1.04 | 1.04
| Std Dev. |  | 0.0042426 | 17
| Rel.Std.Dev. |  | 0.4079%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 241b (Indicative Value 3.24 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 3.22 | 1
| 1 | 2 | 3.25 | 14
| Mean |  | 3.235 | 16
| Median |  | 3.235 | 3.235
| Std Dev. |  | 0.021213 | 20
| Rel.Std.Dev. |  | 0.6557%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 241b (Indicative Value 0.17 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.17 | 1
| 1 | 2 | 0.17 | 16
| Mean |  | 0.17 | 16
| Median |  | 0.17 | 0.17
| Std Dev. |  | 0 | 22
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 241b (Indicative Value 0.465 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.47 | 1
| 1 | 2 | 0.46 | 16
| Mean |  | 0.465 | 16
| Median |  | 0.465 | 0.465
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 1.521%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 241b (Indicative Value 1.8 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.7 | 1
| 1 | 2 | 1.9 | 18
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0.14142 | 24
| Rel.Std.Dev. |  | 7.857%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 241b (Indicative Value 85 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84.2 | 1
| 1 | 2 | 86.4 | 19
| Mean |  | 85.3 | 16
| Median |  | 85.3 | 85.3
| Std Dev. |  | 1.5556 | 25
| Rel.Std.Dev. |  | 1.824%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 241b (Indicative Value 250 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 246 | 1
| 1 | 2 | 253 | 20
| Mean |  | 249.5 | 16
| Median |  | 249.5 | 249.5
| Std Dev. |  | 4.9497 | 26
| Rel.Std.Dev. |  | 1.984%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 241b (Indicative Value 0.6 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.6 | 1
| 1 | 2 | 0.6 | 21
| Mean |  | 0.6 | 16
| Median |  | 0.6 | 0.6
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 241b (Indicative Value 0.06 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.06 | 22
| Mean |  | 0.06 | 16
| Median |  | 0.06 | 0.06
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 241b (Indicative Value 0.8 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.9 | 23
| Mean |  | 0.8 | 16
| Median |  | 0.8 | 0.8
| Std Dev. |  | 0.14142 | 29
| Rel.Std.Dev. |  | 17.68%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 241b (Indicative Value 14.3 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 14.2 | 1
| 1 | 2 | 14.4 | 24
| Mean |  | 14.3 | 16
| Median |  | 14.3 | 14.3
| Std Dev. |  | 0.14142 | 30
| Rel.Std.Dev. |  | 0.989%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 241b (Indicative Value 41.7 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 41.4 | 1
| 1 | 2 | 41.9 | 25
| Mean |  | 41.65 | 16
| Median |  | 41.65 | 41.65
| Std Dev. |  | 0.35355 | 31
| Rel.Std.Dev. |  | 0.8489%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 241b (Indicative Value 136 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 135 | 1
| 1 | 2 | 137 | 26
| Mean |  | 136 | 16
| Median |  | 136 | 136
| Std Dev. |  | 1.4142 | 32
| Rel.Std.Dev. |  | 1.04%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 241b (Indicative Value 1.17 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.21 | 1
| 1 | 2 | 1.12 | 27
| Mean |  | 1.165 | 16
| Median |  | 1.165 | 1.165
| Std Dev. |  | 0.06364 | 33
| Rel.Std.Dev. |  | 5.463%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 241b (Indicative Value 169 ppm) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 166 | 1
| 1 | 2 | 172 | 28
| Mean |  | 169 | 16
| Median |  | 169 | 169
| Std Dev. |  | 4.2426 | 34
| Rel.Std.Dev. |  | 2.51%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 241b (Indicative Value 3.76 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.64 | 1
| 1 | 2 | 3.88 | 29
| Mean |  | 3.76 | 16
| Median |  | 3.76 | 3.76
| Std Dev. |  | 0.16971 | 35
| Rel.Std.Dev. |  | 4.513%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 241b (Indicative Value 2.41 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.37 | 1
| 1 | 2 | 2.44 | 30
| Mean |  | 2.405 | 16
| Median |  | 2.405 | 2.405
| Std Dev. |  | 0.049497 | 36
| Rel.Std.Dev. |  | 2.058%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 241b (Indicative Value 0.97 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.93 | 1
| 1 | 2 | 1.01 | 31
| Mean |  | 0.97 | 16
| Median |  | 0.97 | 0.97
| Std Dev. |  | 0.056569 | 37
| Rel.Std.Dev. |  | 5.832%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 241b (Indicative Value 14.7 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15 | 1
| 1 | 2 | 14.4 | 32
| Mean |  | 14.7 | 16
| Median |  | 14.7 | 14.7
| Std Dev. |  | 0.42426 | 38
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 241b (Indicative Value 3.29 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.22 | 1
| 1 | 2 | 3.35 | 33
| Mean |  | 3.285 | 16
| Median |  | 3.285 | 3.285
| Std Dev. |  | 0.091924 | 39
| Rel.Std.Dev. |  | 2.798%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 241b (Indicative Value 1.55 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.55 | 1
| 1 | 2 | 1.55 | 34
| Mean |  | 1.55 | 16
| Median |  | 1.55 | 1.55
| Std Dev. |  | 0 | 40
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 241b (Indicative Value 2.15 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.15 | 1
| 1 | 2 | 2.14 | 18
| Mean |  | 2.145 | 16
| Median |  | 2.145 | 2.145
| Std Dev. |  | 0.0070711 | 24
| Rel.Std.Dev. |  | 0.3297%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 241b (Indicative Value 0.83 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.83 | 1
| 1 | 2 | 0.82 | 19
| Mean |  | 0.825 | 16
| Median |  | 0.825 | 0.825
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.8571%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 241b (Indicative Value 0.075 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.1 | 20
| Mean |  | 0.075 | 16
| Median |  | 0.075 | 0.075
| Std Dev. |  | 0.035355 | 26
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 241b (Indicative Value 6.68 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.6 | 1
| 1 | 2 | 6.76 | 21
| Mean |  | 6.68 | 16
| Median |  | 6.68 | 6.68
| Std Dev. |  | 0.11314 | 27
| Rel.Std.Dev. |  | 1.694%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 241b (Indicative Value 0.33 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.31 | 1
| 1 | 2 | 0.34 | 22
| Mean |  | 0.325 | 16
| Median |  | 0.325 | 0.325
| Std Dev. |  | 0.021213 | 28
| Rel.Std.Dev. |  | 6.527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 241b (Indicative Value 0.141 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.14 | 1
| 1 | 2 | 0.141 | 23
| Mean |  | 0.1405 | 16
| Median |  | 0.1405 | 0.1405
| Std Dev. |  | 0.00070711 | 29
| Rel.Std.Dev. |  | 0.5033%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 241b (Indicative Value 2.2 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.2 | 1
| 1 | 2 | 2.2 | 24
| Mean |  | 2.2 | 16
| Median |  | 2.2 | 2.2
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 241b (Indicative Value 3.79 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.75 | 1
| 1 | 2 | 3.83 | 7
| Mean |  | 3.79 | 16
| Median |  | 3.79 | 3.79
| Std Dev. |  | 0.056569 | 31
| Rel.Std.Dev. |  | 1.493%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 241b (Indicative Value 9.32 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.51 | 1
| 1 | 2 | 9.12 | 26
| Mean |  | 9.315 | 16
| Median |  | 9.315 | 9.315
| Std Dev. |  | 0.27577 | 32
| Rel.Std.Dev. |  | 2.961%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 241b (Indicative Value 87 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 82 | 1
| 1 | 2 | 92 | 27
| Mean |  | 87 | 16
| Median |  | 87 | 87
| Std Dev. |  | 7.0711 | 33
| Rel.Std.Dev. |  | 8.128%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 241b (Indicative Value 38 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38 | 1
| 1 | 2 | 38 | 28
| Mean |  | 38 | 16
| Median |  | 38 | 38
| Std Dev. |  | 0 | 34
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 241b (Indicative Value 2.02 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.02 | 1
| 1 | 2 | 2.02 | 29
| Mean |  | 2.02 | 16
| Median |  | 2.02 | 2.02
| Std Dev. |  | 0 | 35
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 241b (Indicative Value 17.6 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.5 | 1
| 1 | 2 | 17.7 | 30
| Mean |  | 17.6 | 16
| Median |  | 17.6 | 17.6
| Std Dev. |  | 0.14142 | 36
| Rel.Std.Dev. |  | 0.8035%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 241b (Indicative Value 0.015 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.02 | 1
| 1 | 2 | 0.01 | 31
| Mean |  | 0.015 | 16
| Median |  | 0.015 | 0.015
| Std Dev. |  | 0.0070711 | 37
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 241b (Indicative Value 2.1 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2 | 1
| 1 | 2 | 2.2 | 32
| Mean |  | 2.1 | 16
| Median |  | 2.1 | 2.1
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 6.734%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 241b (Indicative Value 36.4 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 36.4 | 1
| 1 | 2 | 36.3 | 33
| Mean |  | 36.35 | 16
| Median |  | 36.35 | 36.35
| Std Dev. |  | 0.070711 | 39
| Rel.Std.Dev. |  | 0.1945%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 241b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 7
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 40
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 241b (Indicative Value 2.66 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.61 | 1
| 1 | 2 | 2.71 | 18
| Mean |  | 2.66 | 16
| Median |  | 2.66 | 2.66
| Std Dev. |  | 0.070711 | 24
| Rel.Std.Dev. |  | 2.658%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 241b (Indicative Value 1.3 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.2 | 1
| 1 | 2 | 1.4 | 19
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0.14142 | 25
| Rel.Std.Dev. |  | 10.88%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 241b (Indicative Value 152 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 151 | 1
| 1 | 2 | 153 | 20
| Mean |  | 152 | 16
| Median |  | 152 | 152
| Std Dev. |  | 1.4142 | 26
| Rel.Std.Dev. |  | 0.9304%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 241b (Indicative Value 0.24 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.23 | 1
| 1 | 2 | 0.25 | 21
| Mean |  | 0.24 | 16
| Median |  | 0.24 | 0.24
| Std Dev. |  | 0.014142 | 27
| Rel.Std.Dev. |  | 5.893%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 241b (Indicative Value 0.55 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.53 | 1
| 1 | 2 | 0.57 | 22
| Mean |  | 0.55 | 16
| Median |  | 0.55 | 0.55
| Std Dev. |  | 0.028284 | 28
| Rel.Std.Dev. |  | 5.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 241b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 23
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 29
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 241b (Indicative Value 1.31 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.29 | 1
| 1 | 2 | 1.33 | 24
| Mean |  | 1.31 | 16
| Median |  | 1.31 | 1.31
| Std Dev. |  | 0.028284 | 30
| Rel.Std.Dev. |  | 2.159%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 241b (Indicative Value 0.618 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.616 | 1
| 1 | 2 | 0.62 | 25
| Mean |  | 0.618 | 16
| Median |  | 0.618 | 0.618
| Std Dev. |  | 0.0028284 | 31
| Rel.Std.Dev. |  | 0.4577%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 241b (Indicative Value 0.2 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.2 | 1
| 1 | 2 | 0.2 | 26
| Mean |  | 0.2 | 16
| Median |  | 0.2 | 0.2
| Std Dev. |  | 0 | 32
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 241b (Indicative Value 0.34 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.34 | 27
| Mean |  | 0.34 | 16
| Median |  | 0.34 | 0.34
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 241b (Indicative Value 0.37 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.36 | 1
| 1 | 2 | 0.38 | 28
| Mean |  | 0.37 | 16
| Median |  | 0.37 | 0.37
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 3.822%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 241b (Indicative Value 263 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 264 | 1
| 1 | 2 | 261 | 29
| Mean |  | 262.5 | 16
| Median |  | 262.5 | 262.5
| Std Dev. |  | 2.1213 | 35
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 241b (Indicative Value 38 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38 | 1
| 1 | 2 | 38 | 30
| Mean |  | 38 | 16
| Median |  | 38 | 38
| Std Dev. |  | 0 | 36
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 241b (Indicative Value 20.6 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.5 | 1
| 1 | 2 | 20.7 | 31
| Mean |  | 20.6 | 16
| Median |  | 20.6 | 20.6
| Std Dev. |  | 0.14142 | 37
| Rel.Std.Dev. |  | 0.6865%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 241b (Indicative Value 2.36 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.33 | 1
| 1 | 2 | 2.38 | 32
| Mean |  | 2.355 | 16
| Median |  | 2.355 | 2.355
| Std Dev. |  | 0.035355 | 38
| Rel.Std.Dev. |  | 1.501%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 241b (Indicative Value 173 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 165 | 1
| 1 | 2 | 180 | 33
| Mean |  | 172.5 | 16
| Median |  | 172.5 | 172.5
| Std Dev. |  | 10.6066 | 39
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 241b (Indicative Value 75 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 73.5 | 1
| 1 | 2 | 75.5 | 34
| Mean |  | 74.5 | 16
| Median |  | 74.5 | 74.5
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 1.898%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.