## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 505b
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppb) | 554.129 | 549.93 | 558.3279 | 549.0264 | 559.2315
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppb) | 543.98 | 534.9935 | 552.9666 | 538.4928 | 549.4672
4-Acid Digestion
Ag, Silver (ppm) | 1.287 | 1.2197 | 1.3543 | 1.2427 | 1.3313
Al, Aluminium (wt.%) | 7.5198 | 7.3152 | 7.7244 | 7.3636 | 7.6761
As, Arsenic (ppm) | 9.4523 | 8.2282 | 10.6764 | 8.9663 | 9.9383
Ba, Barium (ppm) | 954.93 | 922.6393 | 987.2206 | 933.3757 | 976.4842
Be, Beryllium (ppm) | 2.3738 | 2.2192 | 2.5283 | 2.3034 | 2.4442
Bi, Bismuth (ppm) | 0.72036 | 0.60239 | 0.83832 | 0.58948 | 0.85123
Ca, Calcium (wt.%) | 1.8688 | 1.8157 | 1.9219 | 1.8204 | 1.9172
Cd, Cadmium (ppm) | 0.29222 | 0.25931 | 0.32512 | 0.2723 | 0.31213
Ce, Cerium (ppm) | 67.981 | 63.1137 | 72.8483 | 65.469 | 70.493
Co, Cobalt (ppm) | 9.1779 | 8.673 | 9.6827 | 8.8598 | 9.4959
Cr, Chromium (ppm) | 58.1552 | 54.9885 | 61.322 | 55.3916 | 60.9189
Cs, Caesium (ppm) | 9.5392 | 9.1292 | 9.9492 | 9.2597 | 9.8187
Cu, Copper (wt.%) | 0.31974 | 0.31253 | 0.32694 | 0.31324 | 0.32624
Dy, Dysprosium (ppm) | 3.3024 | 2.7673 | 3.8375 | 3.1729 | 3.4318
Er, Erbium (ppm) | 1.3622 | 1.1593 | 1.5652 | 1.2714 | 1.4531
Eu, Europium (ppm) | 1.3608 | 1.2515 | 1.47 | 1.3057 | 1.4159
Fe, Iron (wt.%) | 3.0429 | 2.9504 | 3.1355 | 2.9837 | 3.1022
Ga, Gallium (ppm) | 19.7327 | 18.7332 | 20.7321 | 19.1965 | 20.2689
Gd, Gadolinium (ppm) | 5.1996 | 4.3594 | 6.0398 | 4.9734 | 5.4257
Hf, Hafnium (ppm) | 2.0442 | 1.8964 | 2.1921 | 1.9401 | 2.1484
Ho, Holmium (ppm) | 0.54333 | 0.41488 | 0.67177 | 0.51441 | 0.57224
In, Indium (ppm) | 0.081453 | 0.071275 | 0.091632 | 0.068868 | 0.094039
K, Potassium (wt.%) | 3.1874 | 3.1045 | 3.2703 | 3.105 | 3.2698
La, Lanthanum (ppm) | 33.6182 | 31.8009 | 35.4354 | 32.2983 | 34.938
Li, Lithium (ppm) | 47.0743 | 45.1717 | 48.977 | 45.8467 | 48.3019
Lu, Lutetium (ppm) | 0.18721 | 0.16455 | 0.20987 | IND | IND
Mg, Magnesium (wt.%) | 0.78878 | 0.76762 | 0.80995 | 0.77132 | 0.80625
Mn, Manganese (wt.%) | 0.042298 | 0.041011 | 0.043585 | 0.041402 | 0.043194
Mo, Molybdenum (ppm) | 68.1368 | 65.9552 | 70.3184 | 66.7546 | 69.519
Na, Sodium (wt.%) | 2.1056 | 2.0526 | 2.1586 | 2.0473 | 2.164
Nb, Niobium (ppm) | 11.8887 | 11.0877 | 12.6898 | 11.4787 | 12.2987
Nd, Neodymium (ppm) | 28.3643 | 23.5418 | 33.1868 | 27.2028 | 29.5258
Ni, Nickel (ppm) | 40.1052 | 38.6865 | 41.5238 | 38.9463 | 41.264
P, Phosphorus (wt.%) | 0.085849 | 0.082923 | 0.088776 | 0.084029 | 0.08767
Pb, Lead (ppm) | 49.7165 | 47.5947 | 51.8383 | 48.3298 | 51.1032
Pr, Praseodymium (ppm) | 7.7969 | 7.0852 | 8.5086 | 7.4651 | 8.1286
Rb, Rubidium (ppm) | 152.7374 | 144.7935 | 160.6813 | 148.8965 | 156.5783
Re, Rhenium (ppm) | 0.022515 | 0.01875 | 0.026281 | 0.019423 | 0.025608
S, Sulphur (wt.%) | 0.53693 | 0.51758 | 0.55627 | 0.52624 | 0.54762
Sb, Antimony (ppm) | 10.0879 | 9.5545 | 10.6213 | 9.7516 | 10.4243
Sc, Scandium (ppm) | 8.578 | 8.0716 | 9.0845 | 8.337 | 8.8191
Se, Selenium (ppm) | 3.9828 | 3.2126 | 4.7531 | 3.2372 | 4.7284
Sm, Samarium (ppm) | 6.2598 | 5.5566 | 6.9631 | 5.9641 | 6.5556
Sn, Tin (ppm) | 4.4083 | 4.1125 | 4.7041 | 4.1648 | 4.6517
Sr, Strontium (ppm) | 265.9939 | 256.2396 | 275.7483 | 260.4404 | 271.5475
Ta, Tantalum (ppm) | 1.0465 | 0.99018 | 1.1029 | 0.99956 | 1.0935
Tb, Terbium (ppm) | 0.69477 | 0.59024 | 0.79931 | 0.65463 | 0.73491
Te, Tellurium (ppm) | 0.17249 | 0.13334 | 0.21163 | 0.13292 | 0.21205
Th, Thorium (ppm) | 12.3089 | 11.416 | 13.2018 | 11.8755 | 12.7424
Ti, Titanium (wt.%) | 0.33719 | 0.32798 | 0.34639 | 0.32904 | 0.34533
Tl, Thallium (ppm) | 0.81893 | 0.77363 | 0.86423 | 0.78073 | 0.85713
Tm, Thulium (ppm) | 0.17929 | 0.14754 | 0.21104 | IND | IND
U, Uranium (ppm) | 3.3984 | 2.9383 | 3.8586 | 3.0867 | 3.7102
V, Vanadium (ppm) | 82.0769 | 80.0837 | 84.0701 | 80.359 | 83.7948
W, Tungsten (ppm) | 9.1165 | 8.2898 | 9.9432 | 8.643 | 9.5899
Y, Yttrium (ppm) | 14.8557 | 14.0365 | 15.6749 | 14.3904 | 15.3211
Yb, Ytterbium (ppm) | 1.1706 | 1.0267 | 1.3146 | 1.1204 | 1.2208
Zn, Zinc (ppm) | 98.2669 | 94.8899 | 101.6439 | 95.8052 | 100.7286
Zr, Zirconium (ppm) | 68.2273 | 64.6349 | 71.8197 | 66.3561 | 70.0986
Aqua Regia Digestion
Ag, Silver (ppm) | 1.2838 | 1.2264 | 1.3412 | 1.2333 | 1.3342
Al, Aluminium (wt.%) | 1.9006 | 1.8483 | 1.9528 | 1.8572 | 1.9439
As, Arsenic (ppm) | 9.0064 | 8.2592 | 9.7536 | 8.5889 | 9.4238
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Ba, Barium (ppm) | 453.678 | 437.0226 | 470.3334 | 444.5194 | 462.8366
Be, Beryllium (ppm) | 1.4834 | 1.399 | 1.5677 | 1.4415 | 1.5252
Bi, Bismuth (ppm) | 0.74516 | 0.61535 | 0.87497 | 0.63984 | 0.85048
Ca, Calcium (wt.%) | 0.78532 | 0.75887 | 0.81177 | 0.77036 | 0.80028
Cd, Cadmium (ppm) | 0.19382 | 0.17424 | 0.21339 | 0.1763 | 0.21133
Ce, Cerium (ppm) | 34.0264 | 31.7952 | 36.2576 | 32.935 | 35.1178
Co, Cobalt (ppm) | 8.829 | 8.3871 | 9.2708 | 8.5896 | 9.0684
Cr, Chromium (ppm) | 54.6312 | 52.7128 | 56.5496 | 53.4179 | 55.8446
Cs, Caesium (ppm) | 8.0329 | 7.6416 | 8.4241 | 7.8153 | 8.2505
Cu, Copper (wt.%) | 0.32236 | 0.31349 | 0.33122 | 0.31754 | 0.32717
Fe, Iron (wt.%) | 2.9583 | 2.8889 | 3.0276 | 2.9014 | 3.0151
Ga, Gallium (ppm) | 8.9227 | 8.5128 | 9.3327 | 8.6605 | 9.1849
Ge, Germanium (ppm) | 0.10417 | 0.079871 | 0.12846 | IND | IND
Hf, Hafnium (ppm) | 0.31635 | 0.29074 | 0.34195 | 0.29542 | 0.33727
Hg, Mercury (ppm) | 0.020987 | 0.0072124 | 0.034761 | IND | IND
In, Indium (ppm) | 0.076657 | 0.070288 | 0.083026 | 0.06865 | 0.084664
K, Potassium (wt.%) | 0.91056 | 0.8805 | 0.94063 | 0.89356 | 0.92756
La, Lanthanum (ppm) | 16.5502 | 15.5191 | 17.5814 | 16.0056 | 17.0949
Li, Lithium (ppm) | 37.5627 | 36.098 | 39.0274 | 36.7748 | 38.3506
Lu, Lutetium (ppm) | 0.097603 | 0.078682 | 0.11652 | IND | IND
Mg, Magnesium (wt.%) | 0.73647 | 0.71392 | 0.75901 | 0.72116 | 0.75178
Mn, Manganese (wt.%) | 0.037909 | 0.036913 | 0.038905 | 0.03727 | 0.038548
Mo, Molybdenum (ppm) | 66.9794 | 64.6311 | 69.3276 | 65.7555 | 68.2032
Na, Sodium (wt.%) | 0.15322 | 0.14249 | 0.16396 | 0.14825 | 0.1582
Ni, Nickel (ppm) | 38.5838 | 37.0027 | 40.1649 | 37.6846 | 39.483
P, Phosphorus (wt.%) | 0.067924 | 0.065993 | 0.069855 | 0.066472 | 0.069376
Pb, Lead (ppm) | 34.2693 | 32.9747 | 35.5639 | 33.154 | 35.3845
Rb, Rubidium (ppm) | 89.3042 | 86.1122 | 92.4963 | 86.8248 | 91.7837
Re, Rhenium (ppm) | 0.021811 | 0.019548 | 0.024073 | 0.020138 | 0.023483
S, Sulphur (wt.%) | 0.53674 | 0.51844 | 0.55503 | 0.52473 | 0.54874
Sb, Antimony (ppm) | 7.6143 | 7.1291 | 8.0996 | 7.335 | 7.8937
Sc, Scandium (ppm) | 7.4821 | 7.0849 | 7.8793 | 7.2311 | 7.733
Se, Selenium (ppm) | 3.7153 | 3.3083 | 4.1222 | 3.5093 | 3.9212
Sn, Tin (ppm) | 3.1256 | 2.9043 | 3.3468 | 2.9574 | 3.2938
Sr, Strontium (ppm) | 90.7258 | 87.2535 | 94.1981 | 88.7667 | 92.6849
Ta, Tantalum (ppm) | < 0.01 | IND | IND | IND | IND
Tb, Terbium (ppm) | 0.49519 | 0.40088 | 0.5895 | 0.47253 | 0.51784
Te, Tellurium (ppm) | 0.16436 | 0.14128 | 0.18744 | IND | IND
Th, Thorium (ppm) | 6.4778 | 5.9497 | 7.0059 | 6.1579 | 6.7977
Ti, Titanium (wt.%) | 0.23753 | 0.22909 | 0.24597 | 0.23225 | 0.24281
Tl, Thallium (ppm) | 0.54266 | 0.5137 | 0.57161 | 0.5143 | 0.57102
U, Uranium (ppm) | 2.9645 | 2.5514 | 3.3777 | 2.7673 | 3.1617
V, Vanadium (ppm) | 71.1402 | 69.098 | 73.1823 | 69.7393 | 72.5411
W, Tungsten (ppm) | 5.8129 | 5.2547 | 6.371 | 5.5594 | 6.0663
Y, Yttrium (ppm) | 10.5684 | 10.1629 | 10.9739 | 10.2799 | 10.8569
Yb, Ytterbium (ppm) | 0.73257 | 0.67008 | 0.79506 | IND | IND
Zn, Zinc (ppm) | 94.5246 | 91.1261 | 97.9231 | 92.7667 | 96.2824
Zr, Zirconium (ppm) | 8.5573 | 7.9439 | 9.1708 | 8.2257 | 8.889
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 505b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 23.6667 | Pt | ppb | 3.3333
4-Acid Digestion
B | ppm | 5383.3386 | Hg | ppm | 0.013
Ge | ppm | 0.22983 | Si | wt.% | 37.4025
Aqua Regia Digestion
Dy | ppm | 2.5105 | Nb | ppm | 0.88997 | Si | wt.% | 0.028597
Er | ppm | 0.96548 | Nd | ppm | 15.0472 | Sm | ppm | 3.3838
Eu | ppm | 0.42652 | Pd | ppb | 22.6527 | Tm | ppm | 0.11622
Gd | ppm | 3.4573 | Pr | ppm | 3.7619
Ho | ppm | 0.37017 | Pt | ppb | 1.3333
Borate Fusion XRF
Al2O3 | wt.% | 15.005 | Fe | wt.% | 3.18 | S | wt.% | 0.5545
As | ppm | 10 | K2O | wt.% | 3.955 | SiO2 | wt.% | 66.505
Ba | ppm | 990 | MgO | wt.% | 1.37 | Sn | ppm | 20
CaO | wt.% | 2.615 | MnO | wt.% | 0.058 | Sr | ppm | 340
Cl | ppm | 40 | Na2O | wt.% | 2.925 | TiO2 | wt.% | 0.5865
Co | ppm | 20 | Ni | ppm | 50 | V | ppm | 75
Cr | ppm | 55 | P | wt.% | 0.0865 | Zn | ppm | 115
Cu | wt.% | 0.3285 | Pb | ppm | 70 | Zr | ppm | 235
Thermogravimetry
LOI1000 | wt.% | 1.465
Infrared Combustion
C | wt.% | 0.135 | S | wt.% | 0.47
Laser Ablation ICP-MS
Ag | ppm | 1.3 | Hf | ppm | 6.2 | Sm | ppm | 6.65
As | ppm | 9 | Ho | ppm | 1.025 | Sn | ppm | 4.3
Ba | ppm | 927 | In | ppm | 0.05 | Sr | ppm | 262
Be | ppm | 2.9 | La | ppm | 35.4 | Ta | ppm | 1.04
Bi | ppm | 0.84 | Lu | ppm | 0.355 | Tb | ppm | 0.905
Cd | ppm | 0.35 | Mn | wt.% | 0.0409 | Te | ppm | < 0.2
Ce | ppm | 68.35 | Mo | ppm | 66.2 | Th | ppm | 12.85
Co | ppm | 9.25 | Nb | ppm | 12 | Ti | wt.% | 0.3545
Cr | ppm | 69.5 | Nd | ppm | 32.55 | Tl | ppm | 0.4
Cs | ppm | 9.265 | Ni | ppm | 48 | Tm | ppm | 0.395
Cu | wt.% | 0.3185 | Pb | ppm | 49.5 | U | ppm | 4.19
Dy | ppm | 5.105 | Pr | ppm | 8.59 | V | ppm | 84.55
Er | ppm | 2.84 | Rb | ppm | 148 | W | ppm | 9.75
Eu | ppm | 1.26 | Re | ppm | < 0.01 | Y | ppm | 26.7
Ga | ppm | 18.3 | Sb | ppm | 10.55 | Yb | ppm | 2.535
Gd | ppm | 5.98 | Sc | ppm | 8.75 | Zn | ppm | 97.5
Ge | ppm | 1.35 | Se | ppm | < 5 | Zr | ppm | 220
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 505b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppb | 554.129 | 13.4389 | 527.2511 | 581.0068 | 513.8121 | 594.4458 | 2.425% | 4.85% | 7.276% | 526.4225 | 581.8354
Aqua Regia Digestion (sample weights 10-50g)
Au, ppb | 543.98 | 21.3078 | 501.3644 | 586.5956 | 480.0566 | 607.9034 | 3.917% | 7.834% | 11.75% | 516.781 | 571.179
4-Acid Digestion
Ag, ppm | 1.287 | 0.044167 | 1.1987 | 1.3753 | 1.1545 | 1.4195 | 3.432% | 6.864% | 10.3% | 1.2226 | 1.3513
Al, wt.% | 7.5198 | 0.25166 | 7.0165 | 8.0232 | 6.7648 | 8.2748 | 3.347% | 6.693% | 10.04% | 7.1438 | 7.8958
As, ppm | 9.4523 | 1.6594 | 6.1334 | 12.7712 | 4.474 | 14.4306 | 17.56% | 35.11% | 52.67% | 8.9797 | 9.9249
Ba, ppm | 954.93 | 51.1204 | 852.6892 | 1057.1707 | 801.5688 | 1108.2911 | 5.353% | 10.71% | 16.06% | 907.1835 | 1002.6765
Be, ppm | 2.3738 | 0.17248 | 2.0288 | 2.7187 | 1.8563 | 2.8912 | 7.266% | 14.53% | 21.8% | 2.2551 | 2.4925
Bi, ppm | 0.72036 | 0.093561 | 0.53323 | 0.90748 | 0.43967 | 1.001 | 12.99% | 25.98% | 38.96% | 0.68434 | 0.75637
Ca, wt.% | 1.8688 | 0.047491 | 1.7738 | 1.9638 | 1.7263 | 2.0113 | 2.541% | 5.083% | 7.624% | 1.7754 | 1.9622
Cd, ppm | 0.29222 | 0.022583 | 0.24705 | 0.33738 | 0.22446 | 0.35997 | 7.728% | 15.46% | 23.19% | 0.2776 | 0.30683
Ce, ppm | 67.981 | 6.0006 | 55.9798 | 79.9822 | 49.9792 | 85.9828 | 8.827% | 17.65% | 26.48% | 64.5819 | 71.38
Co, ppm | 9.1779 | 0.6221 | 7.9337 | 10.422 | 7.3116 | 11.0441 | 6.778% | 13.56% | 20.33% | 8.719 | 9.6367
Cr, ppm | 58.1552 | 4.8151 | 48.525 | 67.7855 | 43.7099 | 72.6006 | 8.28% | 16.56% | 24.84% | 55.2475 | 61.063
Cs, ppm | 9.5392 | 0.62164 | 8.2959 | 10.7825 | 7.6743 | 11.4041 | 6.517% | 13.03% | 19.55% | 9.0623 | 10.0162
Cu, wt.% | 0.31974 | 0.0087919 | 0.30215 | 0.33732 | 0.29336 | 0.34611 | 2.75% | 5.499% | 8.249% | 0.30375 | 0.33573
Dy, ppm | 3.3024 | 0.36415 | 2.5741 | 4.0307 | 2.2099 | 4.3948 | 11.03% | 22.05% | 33.08% | 3.1372 | 3.4675
Er, ppm | 1.3622 | 0.14317 | 1.0759 | 1.6486 | 0.93271 | 1.7917 | 10.51% | 21.02% | 31.53% | 1.2941 | 1.4303
Eu, ppm | 1.3608 | 0.076331 | 1.2081 | 1.5134 | 1.1318 | 1.5898 | 5.609% | 11.22% | 16.83% | 1.2927 | 1.4288
Fe, wt.% | 3.0429 | 0.093249 | 2.8564 | 3.2294 | 2.7632 | 3.3227 | 3.064% | 6.129% | 9.193% | 2.8908 | 3.1951
Ga, ppm | 19.7327 | 1.0824 | 17.5678 | 21.8975 | 16.4854 | 22.9799 | 5.485% | 10.97% | 16.46% | 18.746 | 20.7193
Gd, ppm | 5.1996 | 0.60683 | 3.9859 | 6.4133 | 3.3791 | 7.0201 | 11.67% | 23.34% | 35.01% | 4.9396 | 5.4596
Hf, ppm | 2.0442 | 0.16748 | 1.7093 | 2.3792 | 1.5418 | 2.5467 | 8.193% | 16.39% | 24.58% | 1.942 | 2.1465
Ho, ppm | 0.54333 | 0.0911 | 0.36113 | 0.72553 | 0.27003 | 0.81663 | 16.77% | 33.53% | 50.3% | 0.51616 | 0.57049
In, ppm | 0.081453 | 0.0067182 | 0.068017 | 0.094889 | 0.061299 | 0.10161 | 8.248% | 16.5% | 24.74% | 0.07738 | 0.085526
K, wt.% | 3.1874 | 0.084378 | 3.0187 | 3.3562 | 2.9343 | 3.4406 | 2.647% | 5.294% | 7.942% | 3.0281 | 3.3468
La, ppm | 33.6182 | 1.9011 | 29.816 | 37.4203 | 27.915 | 39.3214 | 5.655% | 11.31% | 16.96% | 31.9373 | 35.2991
Li, ppm | 47.0743 | 2.4935 | 42.0874 | 52.0613 | 39.594 | 54.5547 | 5.297% | 10.59% | 15.89% | 44.7206 | 49.4281
Lu, ppm | 0.18721 | 0.010632 | 0.16595 | 0.20847 | 0.15532 | 0.21911 | 5.679% | 11.36% | 17.04% | 0.17785 | 0.19657
Mg, wt.% | 0.78878 | 0.031231 | 0.72632 | 0.85125 | 0.69509 | 0.88248 | 3.959% | 7.919% | 11.88% | 0.74935 | 0.82822
Mn, wt.% | 0.042298 | 0.0011 | 0.040098 | 0.044498 | 0.038998 | 0.045598 | 2.6% | 5.201% | 7.801% | 0.040183 | 0.044413
Mo, ppm | 68.1368 | 2.6321 | 62.8725 | 73.401 | 60.2404 | 76.0331 | 3.863% | 7.726% | 11.59% | 64.7299 | 71.5436
Na, wt.% | 2.1056 | 0.046295 | 2.013 | 2.1982 | 1.9667 | 2.2445 | 2.199% | 4.397% | 6.596% | 2.0003 | 2.2109
Nb, ppm | 11.8887 | 0.7376 | 10.4135 | 13.3639 | 9.6759 | 14.1015 | 6.204% | 12.41% | 18.61% | 11.2943 | 12.4832
Nd, ppm | 28.3643 | 3.9108 | 20.5428 | 36.1858 | 16.632 | 40.0965 | 13.79% | 27.58% | 41.36% | 26.9461 | 29.7825
Ni, ppm | 40.1052 | 1.6575 | 36.7902 | 43.4202 | 35.1327 | 45.0776 | 4.133% | 8.266% | 12.4% | 38.0999 | 42.1104
P, wt.% | 0.085849 | 0.0032316 | 0.079386 | 0.092313 | 0.076154 | 0.095544 | 3.764% | 7.529% | 11.29% | 0.081557 | 0.090142
Pb, ppm | 49.7165 | 2.7362 | 44.2442 | 55.1889 | 41.508 | 57.925 | 5.504% | 11.01% | 16.51% | 47.2307 | 52.2023
Pr, ppm | 7.7969 | 0.51982 | 6.7573 | 8.8365 | 6.2374 | 9.3563 | 6.667% | 13.33% | 20% | 7.407 | 8.1867
Rb, ppm | 152.7374 | 10.0865 | 132.5644 | 172.9103 | 122.4779 | 182.9968 | 6.604% | 13.21% | 19.81% | 145.1005 | 160.3742
Re, ppm | 0.022515 | 0.0023373 | 0.017841 | 0.02719 | 0.015504 | 0.029527 | 10.38% | 20.76% | 31.14% | 0.02139 | 0.023641
S, wt.% | 0.53693 | 0.025365 | 0.4862 | 0.58766 | 0.46083 | 0.61302 | 4.724% | 9.448% | 14.17% | 0.51008 | 0.56378
Sb, ppm | 10.0879 | 0.43247 | 9.223 | 10.9529 | 8.7905 | 11.3853 | 4.287% | 8.574% | 12.86% | 9.5835 | 10.5923
Sc, ppm | 8.578 | 0.50552 | 7.567 | 9.5891 | 7.0615 | 10.0946 | 5.893% | 11.79% | 17.68% | 8.1491 | 9.0069
Se, ppm | 3.9828 | 0.55031 | 2.8822 | 5.0835 | 2.3319 | 5.6338 | 13.82% | 27.63% | 41.45% | 3.7837 | 4.182
Sm, ppm | 6.2598 | 0.52531 | 5.2092 | 7.3105 | 4.6839 | 7.8358 | 8.392% | 16.78% | 25.18% | 5.9469 | 6.5728
Sn, ppm | 4.4083 | 0.37874 | 3.6508 | 5.1657 | 3.2721 | 5.5445 | 8.592% | 17.18% | 25.77% | 4.1879 | 4.6287
Sr, ppm | 265.9939 | 12.8274 | 240.3391 | 291.6487 | 227.5117 | 304.4761 | 4.822% | 9.645% | 14.47% | 252.6942 | 279.2936
Ta, ppm | 1.0465 | 0.050871 | 0.94479 | 1.1483 | 0.89392 | 1.1991 | 4.861% | 9.722% | 14.58% | 0.9942 | 1.0989
Tb, ppm | 0.69477 | 0.079624 | 0.53552 | 0.85402 | 0.4559 | 0.93364 | 11.46% | 22.92% | 34.38% | 0.66003 | 0.72951
Te, ppm | 0.17249 | 0.027894 | 0.1167 | 0.22827 | 0.088804 | 0.25617 | 16.17% | 32.34% | 48.52% | 0.16386 | 0.18111
Th, ppm | 12.3089 | 0.99648 | 10.316 | 14.3019 | 9.3195 | 15.2984 | 8.096% | 16.19% | 24.29% | 11.6935 | 12.9244
Ti, wt.% | 0.33719 | 0.009208 | 0.31877 | 0.3556 | 0.30956 | 0.36481 | 2.731% | 5.462% | 8.192% | 0.32033 | 0.35405
Tl, ppm | 0.81893 | 0.045687 | 0.72756 | 0.9103 | 0.68187 | 0.95599 | 5.579% | 11.16% | 16.74% | 0.77798 | 0.85988
Tm, ppm | 0.17929 | 0.021381 | 0.13653 | 0.22206 | 0.11515 | 0.24344 | 11.93% | 23.85% | 35.78% | 0.17033 | 0.18826
U, ppm | 3.3984 | 0.33924 | 2.72 | 4.0769 | 2.3807 | 4.4162 | 9.982% | 19.96% | 29.95% | 3.2285 | 3.5684
V, ppm | 82.0769 | 2.341 | 77.395 | 86.7589 | 75.054 | 89.0999 | 2.852% | 5.704% | 8.557% | 77.9731 | 86.1808
W, ppm | 9.1165 | 0.81024 | 7.496 | 10.737 | 6.6858 | 11.5472 | 8.888% | 17.78% | 26.66% | 8.6607 | 9.5723
Y, ppm | 14.8557 | 1.0843 | 12.6872 | 17.0243 | 11.6029 | 18.1086 | 7.299% | 14.6% | 21.9% | 14.1129 | 15.5985
Yb, ppm | 1.1706 | 0.12569 | 0.91924 | 1.422 | 0.79355 | 1.5477 | 10.74% | 21.47% | 32.21% | 1.1121 | 1.2292
Zn, ppm | 98.2669 | 4.912 | 88.4429 | 108.0909 | 83.5309 | 113.0029 | 4.999% | 9.997% | 15% | 93.3535 | 103.1802
Zr, ppm | 68.2273 | 4.0004 | 60.2264 | 76.2282 | 56.226 | 80.2286 | 5.863% | 11.73% | 17.59% | 64.8159 | 71.6387
Aqua Regia Digestion
Ag, ppm | 1.2838 | 0.062849 | 1.1581 | 1.4095 | 1.0952 | 1.4723 | 4.896% | 9.791% | 14.69% | 1.2196 | 1.348
Al, wt.% | 1.9006 | 0.077824 | 1.7449 | 2.0562 | 1.6671 | 2.134 | 4.095% | 8.19% | 12.28% | 1.8055 | 1.9956
As, ppm | 9.0064 | 0.91284 | 7.1807 | 10.8321 | 6.2678 | 11.7449 | 10.14% | 20.27% | 30.41% | 8.556 | 9.4567
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ba, ppm | 453.678 | 30.5823 | 392.5133 | 514.8427 | 361.931 | 545.425 | 6.741% | 13.48% | 20.22% | 430.9941 | 476.3619
Be, ppm | 1.4834 | 0.11529 | 1.2528 | 1.7139 | 1.1375 | 1.8292 | 7.772% | 15.54% | 23.32% | 1.4092 | 1.5575
Bi, ppm | 0.74516 | 0.094866 | 0.55543 | 0.93489 | 0.46056 | 1.0298 | 12.73% | 25.46% | 38.19% | 0.7079 | 0.78242
Ca, wt.% | 0.78532 | 0.045601 | 0.69412 | 0.87652 | 0.64852 | 0.92212 | 5.807% | 11.61% | 17.42% | 0.74605 | 0.82459
Cd, ppm | 0.19382 | 0.014067 | 0.16568 | 0.22195 | 0.15162 | 0.23602 | 7.258% | 14.52% | 21.77% | 0.18412 | 0.20351
Ce, ppm | 34.0264 | 2.5899 | 28.8467 | 39.2061 | 26.2569 | 41.796 | 7.611% | 15.22% | 22.83% | 32.3251 | 35.7278
Co, ppm | 8.829 | 0.47117 | 7.8866 | 9.7713 | 7.4155 | 10.2425 | 5.337% | 10.67% | 16.01% | 8.3875 | 9.2704
Cr, ppm | 54.6312 | 2.4069 | 49.8175 | 59.445 | 47.4106 | 61.8519 | 4.406% | 8.811% | 13.22% | 51.8997 | 57.3628
Cs, ppm | 8.0329 | 0.4285 | 7.1759 | 8.8899 | 6.7474 | 9.3184 | 5.334% | 10.67% | 16% | 7.6312 | 8.4345
Cu, wt.% | 0.32236 | 0.0095281 | 0.3033 | 0.34141 | 0.29377 | 0.35094 | 2.956% | 5.912% | 8.867% | 0.30624 | 0.33847
Fe, wt.% | 2.9583 | 0.10405 | 2.7501 | 3.1664 | 2.6461 | 3.2704 | 3.517% | 7.035% | 10.55% | 2.8103 | 3.1062
Ga, ppm | 8.9227 | 0.46378 | 7.9951 | 9.8503 | 7.5314 | 10.3141 | 5.198% | 10.4% | 15.59% | 8.4766 | 9.3688
Ge, ppm | 0.10417 | 0.020404 | 0.063359 | 0.14497 | 0.042955 | 0.16538 | 19.59% | 39.18% | 58.76% | 0.098958 | 0.10938
Hf, ppm | 0.31635 | 0.022145 | 0.27206 | 0.36064 | 0.24991 | 0.38278 | 7% | 14% | 21% | 0.30053 | 0.33217
Hg, ppm | 0.020987 | 0.0070846 | 0.0068175 | 0.035156 | 0 | 0.04224 | 33.76% | 67.51% | 101.3% | 0.019937 | 0.022036
In, ppm | 0.076657 | 0.0039202 | 0.068816 | 0.084497 | 0.064896 | 0.088417 | 5.114% | 10.23% | 15.34% | 0.072824 | 0.080489
K, wt.% | 0.91056 | 0.049704 | 0.81115 | 1.01 | 0.76145 | 1.0597 | 5.459% | 10.92% | 16.38% | 0.86503 | 0.95609
La, ppm | 16.5502 | 1.2769 | 13.9964 | 19.1041 | 12.7194 | 20.3811 | 7.716% | 15.43% | 23.15% | 15.7227 | 17.3777
Li, ppm | 37.5627 | 2.223 | 33.1167 | 42.0087 | 30.8937 | 44.2317 | 5.918% | 11.84% | 17.75% | 35.6846 | 39.4408
Lu, ppm | 0.097603 | 0.012281 | 0.07304 | 0.12217 | 0.060758 | 0.13445 | 12.58% | 25.17% | 37.75% | 0.092723 | 0.10248
Mg, wt.% | 0.73647 | 0.025864 | 0.68474 | 0.7882 | 0.65887 | 0.81406 | 3.512% | 7.024% | 10.54% | 0.69964 | 0.77329
Mn, wt.% | 0.037909 | 0.0010813 | 0.035746 | 0.040072 | 0.034665 | 0.041153 | 2.852% | 5.705% | 8.557% | 0.036014 | 0.039804
Mo, ppm | 66.9794 | 3.2291 | 60.5211 | 73.4376 | 57.292 | 76.6667 | 4.821% | 9.642% | 14.46% | 63.6304 | 70.3283
Na, wt.% | 0.15322 | 0.016497 | 0.12023 | 0.18622 | 0.10374 | 0.20271 | 10.77% | 21.53% | 32.3% | 0.14556 | 0.16089
Ni, ppm | 38.5838 | 1.5045 | 35.5749 | 41.5927 | 34.0704 | 43.0972 | 3.899% | 7.798% | 11.7% | 36.6546 | 40.513
P, wt.% | 0.067924 | 0.0028609 | 0.062203 | 0.073646 | 0.059342 | 0.076507 | 4.212% | 8.424% | 12.64% | 0.064528 | 0.071321
Pb, ppm | 34.2693 | 2.1431 | 29.9831 | 38.5554 | 27.8401 | 40.6985 | 6.254% | 12.51% | 18.76% | 32.5558 | 35.9827
Rb, ppm | 89.3042 | 1.8579 | 85.5885 | 93.02 | 83.7306 | 94.8779 | 2.08% | 4.161% | 6.241% | 84.839 | 93.7695
Re, ppm | 0.021811 | 0.0012243 | 0.019362 | 0.024259 | 0.018138 | 0.025484 | 5.614% | 11.23% | 16.84% | 0.02072 | 0.022901
S, wt.% | 0.53674 | 0.02617 | 0.4844 | 0.58908 | 0.45823 | 0.61525 | 4.876% | 9.752% | 14.63% | 0.5099 | 0.56357
Sb, ppm | 7.6143 | 0.77204 | 6.0703 | 9.1584 | 5.2982 | 9.9305 | 10.14% | 20.28% | 30.42% | 7.2336 | 7.9951
Sc, ppm | 7.4821 | 0.55674 | 6.3686 | 8.5955 | 5.8119 | 9.1523 | 7.441% | 14.88% | 22.32% | 7.108 | 7.8562
Se, ppm | 3.7153 | 0.24479 | 3.2257 | 4.2048 | 2.9809 | 4.4496 | 6.589% | 13.18% | 19.77% | 3.5295 | 3.901
Sn, ppm | 3.1256 | 0.21542 | 2.6947 | 3.5564 | 2.4793 | 3.7718 | 6.892% | 13.78% | 20.68% | 2.9693 | 3.2819
Sr, ppm | 90.7258 | 4.6259 | 81.4741 | 99.9775 | 76.8482 | 104.6033 | 5.099% | 10.2% | 15.3% | 86.1895 | 95.2621
Ta, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tb, ppm | 0.49519 | 0.065303 | 0.36458 | 0.62579 | 0.29928 | 0.6911 | 13.19% | 26.38% | 39.56% | 0.47043 | 0.51995
Te, ppm | 0.16436 | 0.0204 | 0.12356 | 0.20516 | 0.10316 | 0.22556 | 12.41% | 24.82% | 37.24% | 0.15614 | 0.17258
Th, ppm | 6.4778 | 0.8584 | 4.761 | 8.1946 | 3.9026 | 9.053 | 13.25% | 26.5% | 39.75% | 6.1539 | 6.8017
Ti, wt.% | 0.23753 | 0.01299 | 0.21155 | 0.26351 | 0.19856 | 0.2765 | 5.469% | 10.94% | 16.41% | 0.22565 | 0.24941
Tl, ppm | 0.54266 | 0.032618 | 0.47742 | 0.6079 | 0.44481 | 0.64051 | 6.011% | 12.02% | 18.03% | 0.51553 | 0.56979
U, ppm | 2.9645 | 0.27326 | 2.418 | 3.511 | 2.1447 | 3.7843 | 9.218% | 18.44% | 27.65% | 2.8163 | 3.1127
V, ppm | 71.1402 | 2.2622 | 66.6158 | 75.6645 | 64.3536 | 77.9267 | 3.18% | 6.36% | 9.54% | 67.5832 | 74.6972
W, ppm | 5.8129 | 0.76368 | 4.2855 | 7.3402 | 3.5218 | 8.1039 | 13.14% | 26.28% | 39.41% | 5.5222 | 6.1035
Y, ppm | 10.5684 | 0.41896 | 9.7305 | 11.4064 | 9.3115 | 11.8253 | 3.964% | 7.929% | 11.89% | 10.04 | 11.0968
Yb, ppm | 0.73257 | 0.044037 | 0.6445 | 0.82065 | 0.60046 | 0.86468 | 6.011% | 12.02% | 18.03% | 0.69594 | 0.7692
Zn, ppm | 94.5246 | 2.535 | 89.4546 | 99.5945 | 86.9196 | 102.1295 | 2.682% | 5.364% | 8.045% | 89.7983 | 99.2508
Zr, ppm | 8.5573 | 0.81959 | 6.9181 | 10.1965 | 6.0986 | 11.0161 | 9.578% | 19.16% | 28.73% | 8.1295 | 8.9852
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 505b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*MS | fire assay with inductively coupled plasma mass spectroscopy
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
Text Values:
NSS | No Sufficient Sample (Lab 1.21)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 505b
Lab Slot | Laboratory Name and Location
AGAT Laboratories, Mississauga, Ontario, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Brisbane, QLD, Australia
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Tarkwa, Tarkwa, Ghana
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
Reminex Centre de Recherche, Marrakesh, Marrakesh-Safi, Morocco
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
SGS Tarkwa, Tarkwa, Western Region, Ghana
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 505b (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppb) | Upscaled Value (ppb)
2 | 1 | ANSLu | 1 | 8 | 8 | 242148 | 0.082708 | 577.5662 | 580.7819
2 | 1 | ANSLu | 1 | 14 | 4 | 242149 | 0.088024 | 611.9302 | 582.6185
2 | 1 | ANSLu | 1 | 16 | 6 | 242150 | 0.084059 | 618.0025 | 582.943
2 | 1 | ANSLu | 1 | 13 | 3 | 242151 | 0.088145 | 556.6299 | 579.663
2 | 1 | ANSLu | 1 | 15 | 5 | 242152 | 0.085643 | 588.0886 | 581.3443
2 | 1 | ANSLu | 1 | 3 | 3 | 242153 | 0.087687 | 560.2431 | 579.8561
2 | 1 | ANSLu | 1 | 11 | 1 | 242154 | 0.086268 | 597.1004 | 581.8259
2 | 1 | ANSLu | 1 | 19 | 9 | 242155 | 0.082203 | 594.3328 | 581.678
2 | 1 | ANSLu | 1 | 1 | 1 | 242156 | 0.083943 | 544.6018 | 579.0202
2 | 1 | ANSLu | 1 | 10 | 10 | 242157 | 0.087446 | 559.3573 | 579.8088
2 | 1 | ANSLu | 1 | 2 | 2 | 242158 | 0.086318 | 561.0742 | 579.9005
2 | 1 | ANSLu | 1 | 12 | 2 | 242159 | 0.085766 | 566.8831 | 580.211
2 | 1 | ANSLu | 1 | 6 | 6 | 242160 | 0.086025 | 530.9652 | 578.2914
2 | 1 | ANSLu | 1 | 9 | 9 | 242161 | 0.083441 | 568.3623 | 580.2901
2 | 1 | ANSLu | 1 | 20 | 10 | 242162 | 0.08567 | 566.5065 | 580.1909
2 | 1 | ANSLu | 1 | 7 | 7 | 242163 | 0.084656 | 678.9797 | 586.2019
2 | 1 | ANSLu | 1 | 5 | 5 | 242164 | 0.087585 | 575.4529 | 580.669
2 | 1 | ANSLu | 1 | 4 | 4 | 242165 | 0.085124 | 581.9548 | 581.0165
2 | 1 | ANSLu | 1 | 18 | 8 | 242166 | 0.086262 | 593.426 | 581.6296
2 | 1 | ANSLu | 1 | 17 | 7 | 242167 | 0.086777 | 587.8128 | 581.3296
Mean | [6 empty] | 0.085687 | 580.9635 | 580.9635
Median | [7 empty] | 576.5095 | 580.7255
SD | [7 empty] | 31.5769 | 1.6876
RSD | [7 empty] | 5.435% | 0.2905%
RSD Upscale Factor | [7 empty] | 0.053444
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Au in OREAS 505b (Certified Value 554 ppb) | [30 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.29 | 1.30 | ppb
|  |  | INAA | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*MS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*MS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | 0
|  |  | 0.085g | 50g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 25g | 50g | 30g | 30g | 15g | 50g | 50g | 30g | 30g | 50g | 30g | 30g | 50g | 30g | 15g | 50g | 30g | 30g | 30g | 30g | 1
| 1 | 1 | 544.6018 | 563 | 562 | 571 | 573 | 556 | 554 | 550 | 554 | 512 | 540 | 538.2844 | 548 | 548 | 559 | 541 | 540 | 541 | 562.3538 | 550 | 546 | 533 | 548 | 555 | 558 | 571.9 | 482 | 562 | 559 | 1
| 1 | 2 | 561.0742 | 564 | 553 | 587 | 550 | 564 | 546 | 559 | 553 | 538 | 576 | 529.421 | 548 | 590 | 562 | 535 | 545 | 559 | 558.6398 | 540 | 537 | 541 | 551 | 555 | 563 | 571.5 | 494 | 553 | 560 | #N/A
| 1 | 3 | 560.2431 | 566 | 527 | 591 | 572 | 553 | 552 | 553 | 550 | 533 | 556 | 556.3175 | 542 | 567 | 559 | 544.5 | 554 | 554 | 565.1267 | 540 | 534 | 550 | 552 | 539 | 547 | 571.8 | 540 | 557 | 562 | 16
| 1 | 4 | 581.9548 | 566 | 566 | 592 | 554 | 559 | 548 | 538 | 552 | 547 | 568 | 547.0282 | 559 | 550 | 563 | 557.5 | 547 | 548 | 559.236 | 540 | 507 | 553 | 558 | 560 | 533 | 571.9 | 509 | 567 | 589 | 554.129
| 1 | 5 | 575.4529 | 573 | 569 | 551 | 585 | 548 | 554 | 550 | 553 | 532 | 532 | 533.1139 | 553 | 555 | 556 | 563.5 | 542 | 561 | 565.7438 | 550 | 528 | 547 | 555 | 531 | 567 | 571.6 | 458 | 555 | 560 | 7
| 1 | 6 | 530.9652 | 566 | 560 | 571 | 586 | 548 | 548 | 539 | 559 | 534 | 540 | 538.1316 | 549 | 581 | 560 | 577 | 546 | 539 | 560.5937 | 540 | 524 | 534 | 547 | 541 | 569 | 571.7 | 500 | 555 | 556
|  | 7 | 00: 678.9797
|  | 8 | 00: 577.5662
|  | 9 | 00: 568.3623
|  | 10 | 00: 559.3573
|  | 11 | 00: 597.1004
|  | 12 | 00: 566.8831
|  | 13 | 00: 556.6299
|  | 14 | 00: 611.9302
|  | 15 | 00: 588.0886
|  | 16 | 00: 618.0025
|  | 17 | 00: 587.8128
|  | 18 | 00: 593.426
|  | 19 | 00: 594.3328
|  | 20 | 00: 566.5065
| Mean |  | 580.9635 | 566.3333 | 556.1667 | 577.1667 | 570 | 554.6667 | 550.3333 | 548.1667 | 553.5 | 532.6667 | 552 | 540.3828 | 549.8333 | 565.1667 | 559.8333 | 553.0833 | 545.6667 | 550.3333 | 561.949 | 543.3333 | 529.3333 | 543 | 551.8333 | 546.8333 | 556.1667 | 571.7333 | 497.1667 | 558.1667 | 564.3333
| Median |  | 576.5095 | 566 | 561 | 579 | 572.5 | 554.5 | 550 | 550 | 553 | 533.5 | 548 | 538.208 | 548.5 | 561 | 559.5 | 551 | 545.5 | 551 | 561.4737 | 540 | 531 | 544 | 551.5 | 548 | 560.5 | 571.75 | 497 | 556 | 560
| Std Dev. |  | 31.5769 | 3.5024 | 15.3025 | 15.9301 | 15.1658 | 6.314 | 3.4448 | 8.1833 | 3.0166 | 11.5181 | 17.5271 | 9.8035 | 5.7067 | 17.3138 | 2.4833 | 15.7906 | 4.8442 | 9.2014 | 2.9936 | 5.164 | 13.3217 | 8.3666 | 4.1673 | 11.4266 | 13.8046 | 0.1633 | 27.4111 | 5.3072 | 12.242
| Rel.Std.Dev. |  | 5.435% | 0.6184% | 2.751% | 2.76% | 2.661% | 1.138% | 0.6259% | 1.493% | 0.545% | 2.162% | 3.175% | 1.814% | 1.038% | 3.063% | 0.4436% | 2.855% | 0.8878% | 1.672% | 0.5327% | 0.9504% | 2.517% | 1.541% | 0.7552% | 2.09% | 2.482% | 0.02856% | 5.513% | 0.9508% | 2.169%
| PDM3 |  | 4.843% | 2.202% | 0.3677% | 4.157% | 2.864% | 0.09704% | -0.685% | -1.076% | -0.1135% | -3.873% | -0.3842% | -2.481% | -0.7752% | 1.992% | 1.029% | -0.1887% | -1.527% | -0.685% | 1.411% | -1.948% | -4.475% | -2.008% | -0.4143% | -1.317% | 0.3677% | 3.177% | -10.28% | 0.7287% | 1.842%
| Z-Score (Absolute) |  | NA | 1.31 | 0.35 | 2.34 | 1.66 | 0.2 | 0.21 | 0.42 | 0.09 | 1.89 | 0.05 | 1.16 | 0.26 | 1.2 | 0.69 | 0.05 | 0.65 | 0.21 | 0.9 | 0.88 | 2.21 | 0.91 | 0.07 | 0.54 | 0.35 | 1.83 | 5.27 | 0.54 | 1.12
| Analytical results for Pd in OREAS 505b (Indicative Value 23.7 ppb) | [30 empty] | Indicative
Pd | Round | Replicate | 00: Lab | 1
| No | No | 00: 1.23 | ppb
|  |  | 00: FA*OES | 1
|  |  | 00: 30g | 1
| 1 | 1 | 00: 23 | 1
| 1 | 2 | 00: 24 | 1
| 1 | 3 | 00: 23 | 16
| 1 | 4 | 00: 24 | 23.6667
| 1 | 5 | 00: 24 | 7
| 1 | 6 | 00: 24
| Mean |  | 00: 23.6667
| Median |  | 00: 24
| Std Dev. |  | 00: 0.5164
| Rel.Std.Dev. |  | 00: 2.182%
| PDM3 |  | 00: -1.332e-13%
Z-Score (Absolute) | bar: 00
| Analytical results for Pt in OREAS 505b (Indicative Value 3.33 ppb) | [30 empty] | Indicative
Pt | Round | Replicate | 00: Lab | 1
| No | No | 00: 1.23 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 30g | 2
| 1 | 1 | 00: 3 | 1
| 1 | 2 | 00: 3 | 1
| 1 | 3 | 00: 3 | 16
| 1 | 4 | 00: 3 | 3.3333
| 1 | 5 | 00: 4 | 7
| 1 | 6 | 00: 4
| Mean |  | 00: 3.3333
| Median |  | 00: 3
| Std Dev. |  | 00: 0.5164
| Rel.Std.Dev. |  | 00: 15.49%
| PDM3 |  | 00: 1.11e-13%
Z-Score (Absolute) | bar: 00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 505b (Certified Value 544 ppb) | [23 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.14 | 1.15 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.29 | ppb
|  |  | INAA | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*OES/MS | AR*MS | 0
|  |  | 0.085g | 30g | 15g | 15g | 25g | 25g | 25g | 25g | 25g | 25g | 50g | 30g | 50g | 25g | 25g | 25g | 30g | 30g | 50g | 50g | 10g | 25g | 1
| 1 | 1 | 544.6018 | 530 | 549 | 518 | 513 | 545 | 501 | 558 | 556 | 500 | 582.9 | 558 | 543 | 542.005 | 539 | 544 | 524 | 554.1 | 470 | 549 | 548.175 | 549 | 1
| 1 | 2 | 561.0742 | 530 | 565 | 509.3 | 526 | 557 | 504 | 543 | 555 | 550 | 576.5 | 563 | 531 | 551.0266 | 528 | 545 | 525 | 555.8 | 480 | 548 | 557.1 | 545 | #N/A
| 1 | 3 | 560.2431 | 520 | 555 | 514.7 | 520 | 583 | 491 | 530 | 565 | 500 | 572.1 | 564 | 556 | 544.1183 | 529 | 573 | 529 | 556 | 490 | 539 | 558.1 | 549 | 16
| 1 | 4 | 581.9548 | 520 | 539 | 509.5 | 515 | 563 | 538 | 536 | 563 | 530 | 579.4 | 564 | 568 | 560.506 | 528 | 572 | 524 | 557 | 480 | 532 | 561.95 | 561 | 543.98
| 1 | 5 | 575.4529 | 530 | 571 | 500 | 543 | 560 | 527 | 555 | 535 | 500 | 564.4 | 576 | 567 | 541.3555 | 537 | 573 | 524 | 559 | 470 | 564 | 556.6 | 541 | 9
| 1 | 6 | 530.9652 | 530 | 549 | 497.7 | 526 | 563 | 514 | 556 | 574 | 520 | 564.4 | 566 | 558 | 559.7913 | 523 | NSS | 525 | 559 | 490 | 560 | 556.675 | 553
|  | 7 | 00: 678.9797
|  | 8 | 00: 577.5662
|  | 9 | 00: 568.3623
|  | 10 | 00: 559.3573
|  | 11 | 00: 597.1004
|  | 12 | 00: 566.8831
|  | 13 | 00: 556.6299
|  | 14 | 00: 611.9302
|  | 15 | 00: 588.0886
|  | 16 | 00: 618.0025
|  | 17 | 00: 587.8128
|  | 18 | 00: 593.426
|  | 19 | 00: 594.3328
|  | 20 | 00: 566.5065
| Mean |  | 580.9635 | 526.6667 | 554.6667 | 508.2 | 523.8333 | 561.8333 | 512.5 | 546.3333 | 558 | 516.6667 | 573.2833 | 565.1667 | 553.8333 | 549.8005 | 530.6667 | 561.4 | 525.1667 | 556.8167 | 480 | 548.6667 | 556.4333 | 549.6667
| Median |  | 576.5095 | 530 | 552 | 509.4 | 523 | 561.5 | 509 | 549 | 559.5 | 510 | 574.3 | 564 | 557 | 547.5725 | 528.5 | 572 | 524.5 | 556.5 | 480 | 548.5 | 556.8875 | 549
| Std Dev. |  | 31.5769 | 5.164 | 11.6905 | 7.9835 | 10.8336 | 12.3356 | 17.49 | 11.7417 | 13.2061 | 20.6559 | 7.7391 | 5.947 | 14.3863 | 8.7212 | 6.0882 | 15.437 | 1.9408 | 1.9312 | 8.9443 | 12.1271 | 4.5154 | 6.8896
| Rel.Std.Dev. |  | 5.435% | 0.9805% | 2.108% | 1.571% | 2.068% | 2.196% | 3.413% | 2.149% | 2.367% | 3.998% | 1.35% | 1.052% | 2.598% | 1.586% | 1.147% | 2.75% | 0.3696% | 0.3468% | 1.863% | 2.21% | 0.8115% | 1.253%
| PDM3 |  | 6.799% | -3.183% | 1.965% | -6.577% | -3.704% | 3.282% | -5.787% | 0.4326% | 2.577% | -5.021% | 5.387% | 3.895% | 1.811% | 1.07% | -2.447% | 3.202% | -3.458% | 2.36% | -11.76% | 0.8615% | 2.289% | 1.045%
| Z-Score (Absolute) |  | NA | 1.27 | 0.28 | 2.3 | 1.43 | 0.67 | 2.06 | 0.18 | 0.46 | 1.83 | 1.31 | 0.86 | 0.23 | 0.01 | 1.05 | 0.65 | 1.36 | 0.4 | 3.86 | 0.06 | 0.38 | 0
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 505b (Certified Value 1.29 ppm) | [28 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 1.3 | 1.3 | 1.161 | 1.34 | 1.23 | 1.29 | 1.3 | 1.28 | 1.29 | 1.3209 | 1.31 | 0.32 | 1.21 | 1.33 | 1.2469 | 1.25 | 1.26 | 1.31 | 1.31 | 1.241 | < 5 | 1.1018 | 1.2 | 1.38 | < 2 |  |  | 1
| 1 | 2 | 1.3 | 1.3 | 1.206 | 1.29 | 1.26 | 1.33 | 1.34 | 1.26 | 1.25 | 1.3037 | 1.27 | < 0.2 | 1.28 | 1.25 | 1.251 | 1.26 | 1.22 | 1.32 | 1.3 | 1.281 | < 5 | 1.0949 | 1.2 | 1.32 | < 2 |  |  | 25
| 1 | 3 | 1.3 | 1.3 | 1.181 | 1.33 | 1.33 | 1.26 | 1.36 | 1.32 | 1.28 | 1.3086 | 1.34 | 0.569 | 1.22 | 1.33 | 1.2321 | 1.2 | 1.19 | 1.29 | 1.3 | 1.241 | < 5 | 1.1486 | 1.3 | 1.26 | < 2 |  |  | 16
| 1 | 4 | 1.3 | 1.2 | 1.156 | 1.33 | 1.22 | 1.28 | 1.36 | 1.28 | 1.28 | 1.3323 | 1.31 | 0.292 | 1.29 | 1.32 | 1.2832 | 1.21 | 1.3 | 1.33 | 1.34 | 1.251 | < 5 | 1.1202 | 1.2 | 1.37 | < 2 |  |  | 1.287
| 1 | 5 | 1.3 | 1.3 | 1.187 | 1.27 | 1.33 | 1.3 | 1.3 | 1.24 | 1.33 | 1.343 | 1.32 | 0.326 | 1.26 | 1.25 | 1.2707 | 1.2 | 1.24 | 1.31 | 1.31 | 1.261 | < 5 | 1.0897 | 1.3 | 1.39 | < 2 |  |  | 11
| 1 | 6 | 1.3 | 1.3 | 1.195 | 1.26 | 1.2 | 1.25 | 1.37 | 1.34 | 1.27 | 1.326 | 1.29 | < 0.2 | 1.31 | 1.3 | 1.2694 | 1.19 | 1.28 | 1.29 | 1.32 | 1.261 | < 5 | 1.121 | 1.2 | 1.26 | < 2
| Mean |  | 1.3 | 1.2833 | 1.181 | 1.3033 | 1.2617 | 1.285 | 1.3383 | 1.2867 | 1.2833 | 1.3224 | 1.3067 | 0.37675 | 1.2617 | 1.2967 | 1.2589 | 1.2183 | 1.2483 | 1.3083 | 1.3133 | 1.256 |  | 1.1127 | 1.2333 | 1.33
| Median |  | 1.3 | 1.3 | 1.184 | 1.31 | 1.245 | 1.285 | 1.35 | 1.28 | 1.28 | 1.3234 | 1.31 | 0.323 | 1.27 | 1.31 | 1.2602 | 1.205 | 1.25 | 1.31 | 1.31 | 1.256 |  | 1.111 | 1.2 | 1.345
| Std Dev. |  | 0 | 0.040825 | 0.019401 | 0.034448 | 0.056362 | 0.02881 | 0.031252 | 0.037238 | 0.026583 | 0.014692 | 0.024221 | 0.12902 | 0.039707 | 0.037771 | 0.018796 | 0.029269 | 0.040208 | 0.016021 | 0.015055 | 0.015166 |  | 0.021806 | 0.05164 | 0.05933
| Rel.Std.Dev. |  | 0% | 3.181% | 1.643% | 2.643% | 4.467% | 2.242% | 2.335% | 2.894% | 2.071% | 1.111% | 1.854% | 34.25% | 3.147% | 2.913% | 1.493% | 2.402% | 3.221% | 1.225% | 1.146% | 1.207% |  | 1.96% | 4.187% | 4.461%
| PDM3 |  | 1.01% | -0.2847% | -8.236% | 1.269% | -1.968% | -0.1552% | 3.989% | -0.02572% | -0.2847% | 2.752% | 1.528% | -70.73% | -1.968% | 0.7513% | -2.185% | -5.335% | -3.004% | 1.658% | 2.046% | -2.409% |  | -13.54% | -4.17% | 3.341%
| Z-Score (Absolute) |  | NA | NA | 2.43 | 0.45 | 0.53 | 0.02 | 1.27 | 0.06 | 0.02 | 0.9 | 0.53 | 23.51 | 0.53 | 0.29 | 0.59 | 1.55 | 0.84 | 0.57 | 0.69 | 0.66 | 28.6 | 4.03 | NA | 1.08 | 6.68
Results from laboratories 1.01, 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 505b (Certified Value 7.52 wt.%) | [28 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 7.83 | 7.51 | 6.88 | 7.76 | 7.2 | 7.13 | 7.42 | 7.46 | 7.6222 | 7.4148 | 7.83 | 7.52 | 7.6109 | 7.55 | 7.67 | 7.7129 | 7.43 | 7.4 | 7.37 | 6.51 | 8.32 | 7.6067 | 7.3729 | 7.55 | 7.43 | 7.33 |  | 1
| 1 | 2 | 7.91 | 7.06 | 7.11 | 7.69 | 7.36 | 7.41 | 7.57 | 7.63 | 7.7257 | 7.4373 | 7.89 | 7.52 | 7.6189 | 7.56 | 7.6 | 7.6369 | 7.47 | 7.47 | 7.65 | 6.47 | 9.36 | 7.2378 | 7.127 | 7.84 | 7.37 | 7.34 |  | #N/A
| 1 | 3 | 7.96 | 7.13 | 7.14 | 7.72 | 7.38 | 7.27 | 7.47 | 7.49 | 7.8855 | 7.4424 | 7.71 | 7.38 | 7.6624 | 7.53 | 7.54 | 7.5658 | 7.65 | 7.36 | 7.58 | 6.42 | 7.44 | 6.5977 | 7.2698 | 8.06 | 7.4 | 7.38 |  | 16
| 1 | 4 | 7.92 | 7.79 | 7.04 | 7.58 | 7.39 | 7.29 | 7.82 | 7.66 | 7.8764 | 7.3641 | 7.85 | 7.59 | 7.6836 | 7.5 | 7.65 | 7.4926 | 7.55 | 7.39 | 7.73 | 6.44 | 8.26 | 7.3282 | 7.3109 | 7.95 | 7.48 | 7.38 |  | 7.5198
| 1 | 5 | 8.02 | 7.13 | 7.16 | 7.69 | 7.57 | 7.2 | 7.38 | 7.47 | 7.8865 | 7.3928 | 7.74 | 7.42 | 7.6186 | 7.52 | 7.53 | 7.6481 | 7.47 | 7.43 | 7.83 | 6.55 | 7.51 | 7.0703 | 7.1414 | 7.38 | 7.45 | 7.26 |  | 12
| 1 | 6 | 7.93 | 7.33 | 7 | 7.63 | 7.38 | 7.21 | 7.55 | 7.39 | 7.8975 | 7.3585 | 7.93 | 7.63 | 7.6658 | 7.49 | 7.69 | 7.5356 | 7.66 | 7.06 | 7.4 | 6.59 | 7.21 | 7.174 | 7.2604 | 7.88 | 6.88 | 7.3
| Mean |  | 7.9283 | 7.325 | 7.055 | 7.6783 | 7.38 | 7.2517 | 7.535 | 7.5167 | 7.8156 | 7.4016 | 7.825 | 7.51 | 7.6434 | 7.525 | 7.6133 | 7.5986 | 7.5383 | 7.3517 | 7.5933 | 6.4967 | 8.0167 | 7.1691 | 7.2471 | 7.7767 | 7.335 | 7.3317
| Median |  | 7.925 | 7.23 | 7.075 | 7.69 | 7.38 | 7.24 | 7.51 | 7.48 | 7.881 | 7.4038 | 7.84 | 7.52 | 7.6407 | 7.525 | 7.625 | 7.6013 | 7.51 | 7.395 | 7.615 | 6.49 | 7.885 | 7.2059 | 7.2651 | 7.86 | 7.415 | 7.335
| Std Dev. |  | 0.062423 | 0.28169 | 0.10502 | 0.064317 | 0.11747 | 0.096003 | 0.15758 | 0.10539 | 0.11472 | 0.03591 | 0.085264 | 0.095917 | 0.030814 | 0.027386 | 0.067725 | 0.08157 | 0.098472 | 0.1477 | 0.18184 | 0.065625 | 0.79947 | 0.33398 | 0.096109 | 0.25835 | 0.22616 | 0.046655
| Rel.Std.Dev. |  | 0.7873% | 3.846% | 1.489% | 0.8376% | 1.592% | 1.324% | 2.091% | 1.402% | 1.468% | 0.4852% | 1.09% | 1.277% | 0.4031% | 0.3639% | 0.8896% | 1.073% | 1.306% | 2.009% | 2.395% | 1.01% | 9.973% | 4.659% | 1.326% | 3.322% | 3.083% | 0.6363%
| PDM3 |  | 5.432% | -2.591% | -6.181% | 2.108% | -1.859% | -3.566% | 0.2017% | -0.04206% | 3.934% | -1.572% | 4.058% | -0.1307% | 1.643% | 0.06876% | 1.243% | 1.048% | 0.2461% | -2.236% | 0.9775% | -13.61% | 6.607% | -4.664% | -3.627% | 3.415% | -2.458% | -2.502%
| Z-Score (Absolute) |  | 1.55 | 0.74 | 1.77 | 0.6 | 0.54 | 1.02 | 0.05 | 0.02 | 1.12 | 0.45 | 1.16 | 0.04 | 0.47 | 0.02 | 0.35 | 0.3 | 0.07 | 0.64 | 0.28 | 3.89 | 1.88 | 1.34 | 1.04 | 0.97 | 0.71 | 0.72
| Analytical results for As in OREAS 505b (Certified Value 9.45 ppm) | [28 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
2
| 1 | 1 | 10 | 9 | 6.8 | 9 | 10.2 | 9.9 | 9.8 | 8.7 | 10.4 | 13.7436 | 13.5 | 27.93 | 6.11 | 13 | 10.005 | 9.6 | 9.4 | 14.6 | 9.9 | 10.835 | 6.324 | 5.1 | 9 | 12 |  |  |  | 1
| 1 | 2 | 11 | 10 | 6.4 | 8 | 9.2 | 10.4 | 10.2 | 8.7 | 9.3 | 13.8401 | 13.8 | 24.44 | 6.14 | 10 | 9.5665 | 9.8 | 9.2 | 14.5 | 10.1 | 11.602 | 6.32 | 4.9 | 9 | 13 |  |  |  | 13
| 1 | 3 | 11 | 9 | 6 | 9 | 9.3 | 10.4 | 11.4 | 9.5 | 9.8 | 12.9754 | 11.3 | 27.8 | 6.49 | 11 | 10.0526 | 9.1 | 9.5 | 14.4 | 10.5 | 10.893 | 6.202 | 4.4 | 11 | 14 |  |  |  | 16
| 1 | 4 | 11 | 10 | 5.4 | 8 | 9.9 | 9.8 | 10.2 | 9.3 | 10.3 | 13.7662 | 12.5 | 26.41 | 6.31 | 12 | 9.7866 | 8.9 | 9.4 | 14.3 | 10.2 | 10.758 | 6.361 | 4.2 | 10 | 11 |  |  |  | 9.4523
| 1 | 5 | 11 | 9 | 5.2 | 9 | 9.2 | 9.9 | 9.7 | 8.2 | 10.2 | 13.5968 | 12.6 | 26.08 | 6.18 | 11 | 9.9169 | 9.2 | 9.1 | 14.6 | 10 | 10.409 | 6.206 | 5.3 | 10 | 12 |  |  |  | 13
| 1 | 6 | 11 | 10 | 5.8 | 9 | 9.6 | 9.4 | 9.6 | 8.9 | 10.5 | 13.748 | 12.2 | 31.12 | 6.15 | 12 | 9.4773 | 9.4 | 9.6 | 14.4 | 9.6 | 10.367 | 6.329 | 4.8 | 8 | 10
| Mean |  | 10.8333 | 9.5 | 5.9333 | 8.6667 | 9.5667 | 9.9667 | 10.15 | 8.8833 | 10.0833 | 13.6117 | 12.65 | 27.2967 | 6.23 | 11.5 | 9.8008 | 9.3333 | 9.3667 | 14.4667 | 10.05 | 10.8107 | 6.2903 | 4.7833 | 9.5 | 12
| Median |  | 11 | 9.5 | 5.9 | 9 | 9.45 | 9.9 | 10 | 8.8 | 10.25 | 13.7458 | 12.55 | 27.105 | 6.165 | 11.5 | 9.8518 | 9.3 | 9.4 | 14.45 | 10.05 | 10.7965 | 6.322 | 4.85 | 9.5 | 12
| Std Dev. |  | 0.40825 | 0.54772 | 0.60222 | 0.5164 | 0.41312 | 0.38297 | 0.66257 | 0.46655 | 0.4535 | 0.32158 | 0.90499 | 2.2671 | 0.14519 | 1.0488 | 0.23595 | 0.33267 | 0.18619 | 0.12111 | 0.30166 | 0.44599 | 0.068436 | 0.41673 | 1.0488 | 1.4142
| Rel.Std.Dev. |  | 3.768% | 5.766% | 10.15% | 5.958% | 4.318% | 3.843% | 6.528% | 5.252% | 4.498% | 2.363% | 7.154% | 8.306% | 2.33% | 9.12% | 2.407% | 3.564% | 1.988% | 0.8371% | 3.002% | 4.125% | 1.088% | 8.712% | 11.04% | 11.79%
| PDM3 |  | 14.61% | 0.5047% | -37.23% | -8.312% | 1.21% | 5.442% | 7.381% | -6.019% | 6.676% | 44% | 33.83% | 188.8% | -34.09% | 21.66% | 3.687% | -1.259% | -0.9059% | 53.05% | 6.323% | 14.37% | -33.45% | -49.4% | 0.5047% | 26.95%
| Z-Score (Absolute) |  | NA | NA | 2.76 | NA | 0.22 | 0.06 | 0.19 | 0.7 | 0.14 | 2.61 | 1.94 | 12.18 | 2.56 | NA | 0.06 | 0.39 | 0.36 | 3.21 | 0.12 | 0.65 | 2.51 | 3.57 | NA | NA
Results from laboratories 1.01, 1.02, 1.04, 1.17, 1.29 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 505b (Indicative Value 5383 ppm) | [28 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.02: Lab |  |  |  | 1
| No | No | 1.01: 1.10; 1.02: 1.13 |  |  |  | ppm
|  |  | 1.01: 4A*OES; 1.02: 4A*OES/MS |  |  |  | 0
0
| 1 | 1 | 1.01: 8.466; 1.02: 11067.16 |  |  |  | 1
| 1 | 2 | 1.01: 10.4234; 1.02: 9757.03 |  |  |  | 3
| 1 | 3 | 1.01: 10.0993; 1.02: 13096.8 |  |  |  | 16
| 1 | 4 | 1.01: 10.6382; 1.02: 9874.33 |  |  |  | 5383.3386
| 1 | 5 | 1.01: 11.2982; 1.02: 8927.97 |  |  |  | 9
| 1 | 6 | 1.01: 10.6282; 1.02: 11815.22
| Mean |  | 1.01: 10.2589; 1.02: 10756.4183
| Median |  | 1.01: 10.5258; 1.02: 10470.745
| Std Dev. |  | 1.01: 0.96202; 1.02: 1537.211
| Rel.Std.Dev. |  | 1.01: 9.377%; 1.02: 14.29%
| PDM3 |  | 1.01: -99.81%; 1.02: 99.81%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ba in OREAS 505b (Certified Value 955 ppm) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 954 | 870 | 651 | 972 | 910 | 950 | 940 | 940 | 1007 | 1009.9168 | 946 | 1503.29 | 871.89 | 985.9 | 938 | 991.1111 | 969 | 930 | 980 | 838 | 998 | 1027.91 | 850.31 | 879 | 1016 | 905 |  | 1
| 1 | 2 | 947 | 795 | 488 | 965 | 950 | 970 | 970 | 940 | 993.3 | 1057.8627 | 989 | 1312.37 | 904.46 | 977.5 | 937 | 988.1746 | 959 | 950 | 1020 | 845 | 1033 | 1008.5 | 846.461 | 878 | 945 | 908 |  | 14
| 1 | 3 | 963 | 684 | 501 | 963 | 940 | 960 | 960 | 950 | 1004.4 | 1056.0242 | 948 | 1505.81 | 962.09 | 983.2 | 940 | 989.7496 | 951 | 940 | 1010 | 833 | 989 | 998.908 | 846.737 | 909 | 1008 | 912 |  | 16
| 1 | 4 | 964 | 958 | 503 | 950 | 950 | 970 | 1000 | 950 | 1017.7 | 1040.5732 | 961 | 1587.34 | 952.96 | 974.8 | 930 | 977.1244 | 963 | 940 | 1030 | 848 | 1007 | 1016.028 | 845.497 | 884 | 1020 | 912 |  | 954.93
| 1 | 5 | 952 | 690 | 472 | 962 | 970 | 960 | 940 | 920 | 994.9 | 1038.7251 | 950 | 1292.67 | 946.94 | 974.6 | 934 | 981.2316 | 967 | 940 | 1040 | 843 | 991 | 1029.774 | 852.012 | 882 | 1016 | 896 |  | 14
| 1 | 6 | 932 | 737 | 601 | 960 | 940 | 950 | 970 | 940 | 1020 | 1056.8372 | 932 | 1445.83 | 959.31 | 985.4 | 937 | 982.2721 | 961 | 920 | 990 | 836 | 994 | 1020.866 | 853.918 | 892 | 957 | 903
| Mean |  | 952 | 789 | 536 | 962 | 943.3333 | 960 | 963.3333 | 940 | 1006.2167 | 1043.3232 | 954.3333 | 1441.2183 | 932.9417 | 980.2333 | 936 | 984.9439 | 961.6667 | 936.6667 | 1011.6667 | 840.5 | 1002 | 1016.9977 | 849.1558 | 887.3333 | 993.6667 | 906
| Median |  | 953 | 766 | 502 | 962.5 | 945 | 960 | 965 | 940 | 1005.7 | 1048.2987 | 949 | 1474.56 | 949.95 | 980.35 | 937 | 985.2234 | 962 | 940 | 1015 | 840.5 | 996 | 1018.447 | 848.5235 | 883 | 1012 | 906.5
| Std Dev. |  | 11.7813 | 108.3402 | 72.3381 | 7.1833 | 19.6638 | 8.9443 | 22.5093 | 10.9545 | 11.1415 | 18.4393 | 19.3563 | 116.6759 | 36.5516 | 5.2217 | 3.5214 | 5.5431 | 6.4083 | 10.328 | 23.1661 | 5.7533 | 16.4682 | 11.8124 | 3.4254 | 11.7246 | 33.4943 | 6.0992
| Rel.Std.Dev. |  | 1.238% | 13.73% | 13.5% | 0.7467% | 2.085% | 0.9317% | 2.337% | 1.165% | 1.107% | 1.767% | 2.028% | 8.096% | 3.918% | 0.5327% | 0.3762% | 0.5628% | 0.6664% | 1.103% | 2.29% | 0.6845% | 1.644% | 1.161% | 0.4034% | 1.321% | 3.371% | 0.6732%
| PDM3 |  | -0.3068% | -17.38% | -43.87% | 0.7404% | -1.214% | 0.5309% | 0.88% | -1.563% | 5.371% | 9.257% | -0.06248% | 50.92% | -2.303% | 2.65% | -1.982% | 3.143% | 0.7055% | -1.913% | 5.941% | -11.98% | 4.929% | 6.5% | -11.08% | -7.079% | 4.056% | -5.124%
| Z-Score (Absolute) |  | 0.11 | 3.53 | 8.84 | 0.1 | 0.29 | 0.06 | 0.13 | 0.36 | 1.03 | 1.81 | 0.06 | 10.16 | 0.51 | 0.48 | 0.44 | 0.58 | 0.09 | 0.43 | 1.14 | 2.45 | 0.94 | 1.26 | 2.27 | 1.47 | 0.77 | 1.07
| Analytical results for Be in OREAS 505b (Certified Value 2.37 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
3
| 1 | 1 | 2.7 | 3 | 2 | 2.56 | 2.26 | 2.41 | 2.48 | 2.34 | 2.38 | 1.8958 | 2.3192 | 2 | 3.13 | 3 | 2.6242 | 2.42 | 2.3 | 2.42 | < 5 | 3 | < 5 | 2.1146 | 2 | 2.3 | 2.2 |  |  | 1
| 1 | 2 | 2.8 | 3 | 2 | 2.61 | 2.36 | 2.45 | 2.49 | 2.35 | 2.29 | 1.9059 | 2.3398 | 2.13 | 3.14 | 3 | 2.6119 | 2.38 | 2.33 | 2.51 | < 5 | 2 | < 5 | 2.059 | 2 | 2.3 | 2.2 |  |  | 28
| 1 | 3 | 2.6 | 3 | 2 | 2.61 | 2.36 | 2.42 | 2.43 | 2.52 | 2.4 | 1.8662 | 2.3896 | 2.01 | 3.09 | 3 | 2.5764 | 2.24 | 2.33 | 2.48 | < 5 | 2 | < 5 | 2.1027 | 2 | 2.1 | 2.4 |  |  | 16
| 1 | 4 | 2.5 | 3 | 2 | 2.58 | 2.34 | 2.45 | 2.55 | 2.4 | 2.4 | 1.8883 | 2.3084 | 2.03 | 3.09 | 3 | 2.6395 | 2.35 | 2.33 | 2.54 | < 5 | 2 | < 5 | 2.1041 | 2 | 2.2 | 2.4 |  |  | 2.3738
| 1 | 5 | 2.7 | 3 | 3 | 2.43 | 2.4 | 2.43 | 2.44 | 2.31 | 2.54 | 1.9147 | 2.3134 | 2.04 | 3.05 | 3 | 2.5675 | 2.26 | 2.33 | 2.57 | < 5 | 2 | < 5 | 2.046 | 2 | 2.3 | 2.4 |  |  | 15
| 1 | 6 | 2.7 | 3 | 2 | 2.43 | 2.36 | 2.38 | 2.49 | 2.49 | 2.36 | 1.941 | 2.241 | 2.05 | 3.07 | 3 | 2.5801 | 2.3 | 2.27 | 2.43 | < 5 | 2 | < 5 | 2.0937 | 2 | 2.2 | 2.3
| Mean |  | 2.6667 | 3 | 2.1667 | 2.5367 | 2.3467 | 2.4233 | 2.48 | 2.4017 | 2.395 | 1.902 | 2.3186 | 2.0433 | 3.095 | 3 | 2.5999 | 2.325 | 2.315 | 2.4917 |  | 2.1667 |  | 2.0867 | 2 | 2.2333 | 2.3167
| Median |  | 2.7 | 3 | 2 | 2.57 | 2.36 | 2.425 | 2.485 | 2.375 | 2.39 | 1.9008 | 2.3163 | 2.035 | 3.09 | 3 | 2.596 | 2.325 | 2.33 | 2.495 |  | 2 |  | 2.0982 | 2 | 2.25 | 2.35
| Std Dev. |  | 0.10328 | 0 | 0.40825 | 0.084774 | 0.046762 | 0.026583 | 0.042895 | 0.085654 | 0.081915 | 0.025319 | 0.048252 | 0.046332 | 0.034496 | 0 | 0.029296 | 0.070356 | 0.0251 | 0.059805 |  | 0.40825 |  | 0.027622 | 0 | 0.08165 | 0.098319
| Rel.Std.Dev. |  | 3.873% | 0% | 18.84% | 3.342% | 1.993% | 1.097% | 1.73% | 3.566% | 3.42% | 1.331% | 2.081% | 2.267% | 1.115% | 0% | 1.127% | 3.026% | 1.084% | 2.4% |  | 18.84% |  | 1.324% | 0% | 3.656% | 4.244%
| PDM3 |  | 12.34% | 26.38% | -8.724% | 6.863% | -1.141% | 2.088% | 4.475% | 1.176% | 0.8947% | -19.87% | -2.325% | -13.92% | 30.38% | 26.38% | 9.528% | -2.054% | -2.475% | 4.967% |  | -8.724% |  | -12.09% | -15.75% | -5.916% | -2.405%
| Z-Score (Absolute) |  | 1.86 | NA | NA | 0.96 | 0.36 | 0.17 | 0.56 | 0.02 | 0.02 | 3.43 | 0.55 | 2.46 | 4.82 | NA | 1.39 | 0.51 | 0.58 | 0.65 | 0.7 | NA | 0.7 | 2.16 | NA | 1.14 | 0.56
Results from laboratories 1.02, 1.03, 1.17, 1.23 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 505b (Certified Value 0.72 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.7 | 0.7 | 0.65 | 0.66 | 0.85 | 0.73 | 0.63 | 0.58 | 0.83 | 0.90289 | 0.65 | 8.31 | < 5 | 0.77 | 0.72838 | 1.13 | 0.77 | 0.6 | < 5 | 0.68 | < 5 | 0.724 | 0.6 | 0.75 | < 5 |  |  | 1
| 1 | 2 | 0.74 | 0.7 | 0.57 | 0.75 | 0.61 | 0.65 | 0.72 | 0.62 | 0.65 | 0.88972 | 1.27 | 6.95 | < 5 | 0.8 | 0.73557 | 0.94 | 0.85 | 0.68 | < 5 | 0.86 | < 5 | 0.747 | 0.7 | 0.88 | 6 |  |  | 29
| 1 | 3 | 0.67 | 0.7 | 0.68 | 0.64 | 0.66 | 0.78 | 0.73 | 1.04 | 0.69 | 0.80727 | 1.66 | 9.28 | < 5 | 0.6 | 0.73576 | 0.98 | 0.87 | 0.64 | < 5 | 0.63 | < 5 | 0.767 | 0.6 | 0.65 | < 5 |  |  | 16
| 1 | 4 | 0.65 | 0.8 | 0.6 | 0.7 | 0.63 | 0.65 | 1 | 0.6 | 0.74 | 0.78154 | 0.91 | 7.18 | < 5 | 0.76 | 0.68493 | 0.97 | 0.81 | 0.75 | < 5 | 0.7 | < 5 | 0.745 | 0.7 | 0.86 | < 5 |  |  | 0.72036
| 1 | 5 | 0.73 | 0.7 | 0.65 | 0.61 | 0.67 | 0.58 | 0.61 | 0.64 | 0.77 | 0.92402 | 0.58 | 6.98 | < 5 | 0.73 | 0.73502 | 0.88 | 1.22 | 0.62 | < 5 | 0.7 | < 5 | 0.747 | 0.6 | 0.69 | < 5 |  |  | 16
| 1 | 6 | 0.77 | 0.7 | 0.7 | 0.62 | 0.9 | 0.68 | 0.91 | 0.71 | 0.66 | 0.8604 | 0.67 | 7.45 | < 5 | 0.99 | 0.68258 | 0.88 | 0.67 | 0.69 | < 5 | 0.73 | < 5 | 0.734 | 0.8 | 0.67 | < 5
| Mean |  | 0.71 | 0.71667 | 0.64167 | 0.66333 | 0.72 | 0.67833 | 0.76667 | 0.69833 | 0.72333 | 0.86097 | 0.95667 | 7.6917 |  | 0.775 | 0.71704 | 0.96333 | 0.865 | 0.66333 |  | 0.71667 |  | 0.744 | 0.66667 | 0.75 | 6
| Median |  | 0.715 | 0.7 | 0.65 | 0.65 | 0.665 | 0.665 | 0.725 | 0.63 | 0.715 | 0.87506 | 0.79 | 7.315 |  | 0.765 | 0.7317 | 0.955 | 0.83 | 0.66 |  | 0.7 |  | 0.746 | 0.65 | 0.72 | 6
| Std Dev. |  | 0.045166 | 0.040825 | 0.048751 | 0.053166 | 0.12296 | 0.06969 | 0.15603 | 0.17325 | 0.069761 | 0.056128 | 0.42735 | 0.92541 |  | 0.12629 | 0.02594 | 0.092232 | 0.1878 | 0.05465 |  | 0.077632 |  | 0.014478 | 0.08165 | 0.098995
| Rel.Std.Dev. |  | 6.361% | 5.696% | 7.598% | 8.015% | 17.08% | 10.27% | 20.35% | 24.81% | 9.644% | 6.519% | 44.67% | 12.03% |  | 16.3% | 3.618% | 9.574% | 21.71% | 8.239% |  | 10.83% |  | 1.946% | 12.25% | 13.2%
| PDM3 |  | -1.437% | -0.512% | -10.92% | -7.916% | -0.04929% | -5.833% | 6.429% | -3.057% | 0.4134% | 19.52% | 32.8% | 967.8% |  | 7.586% | -0.46% | 33.73% | 20.08% | -7.916% |  | -0.512% |  | 3.282% | -7.453% | 4.115% | 732.9%
| Z-Score (Absolute) |  | 0.33 | NA | 0.93 | 0.74 | 0.24 | 0.61 | 0.17 | 0.43 | 0.21 | 1.01 | 1.86 | 61.48 | 15.52 | 0.25 | 0.27 | 1.92 | 1.04 | 0.74 | 15.52 | 0.27 | 15.52 | 0.03 | NA | 0.03 | NA
Results from laboratories 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution. Results from laboratory 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Ca in OREAS 505b (Certified Value 1.87 wt.%) | [28 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 1.86 | 1.96 | 1.69 | 1.86 | 1.82 | 1.82 | 1.86 | 1.9 | 1.7931 | 1.8374 | 1.85 | 2.28 | 1.9019 | 1.84 | 1.87 | 1.8877 | 1.98 | 1.84 | 1.92 | 1.5 | 1.84 | 1.8754 | 1.953 | 1.84 | 1.65 | 1.813 |  | 1
| 1 | 2 | 1.85 | 1.87 | 1.73 | 1.85 | 1.87 | 1.86 | 1.9 | 1.95 | 1.8164 | 1.8438 | 1.87 | 2.13 | 1.9419 | 1.84 | 1.8 | 1.8785 | 1.99 | 1.85 | 1.99 | 1.46 | 1.9 | 1.8614 | 1.9578 | 1.82 | 1.66 | 1.813 |  | #N/A
| 1 | 3 | 1.85 | 1.93 | 1.74 | 1.86 | 1.88 | 1.83 | 1.86 | 1.86 | 1.8519 | 1.8209 | 1.82 | 2.41 | 1.9724 | 1.83 | 1.81 | 1.842 | 1.85 | 1.84 | 1.97 | 1.53 | 1.81 | 1.7551 | 1.9415 | 1.89 | 1.65 | 1.82 |  | 16
| 1 | 4 | 1.83 | 2.03 | 1.72 | 1.83 | 1.87 | 1.84 | 1.96 | 1.95 | 1.8556 | 1.862 | 1.86 | 2.66 | 1.9556 | 1.86 | 1.86 | 1.8468 | 1.93 | 1.85 | 2.01 | 1.5 | 1.85 | 1.9813 | 1.9409 | 1.88 | 1.67 | 1.82 |  | 1.8688
| 1 | 5 | 1.88 | 1.91 | 1.75 | 1.85 | 1.93 | 1.83 | 1.86 | 1.9 | 1.8569 | 1.8172 | 1.84 | 2.11 | 1.9484 | 1.82 | 1.82 | 1.8666 | 1.93 | 1.86 | 2.04 | 1.55 | 1.85 | 2.7537 | 1.9019 | 1.82 | 1.69 | 1.791 |  | 17
| 1 | 6 | 1.88 | 1.96 | 1.71 | 1.84 | 1.87 | 1.84 | 1.9 | 1.84 | 1.8623 | 1.8292 | 1.87 | 2.25 | 1.943 | 1.89 | 1.84 | 1.8696 | 1.88 | 1.79 | 1.95 | 1.45 | 1.83 | 1.9058 | 1.9298 | 1.87 | 1.57 | 1.8
| Mean |  | 1.8583 | 1.9433 | 1.7233 | 1.8483 | 1.8733 | 1.8367 | 1.89 | 1.9 | 1.8394 | 1.8351 | 1.8517 | 2.3067 | 1.9439 | 1.8467 | 1.8333 | 1.8652 | 1.9267 | 1.8383 | 1.98 | 1.4983 | 1.8467 | 2.0221 | 1.9375 | 1.8533 | 1.6483 | 1.8095
| Median |  | 1.855 | 1.945 | 1.725 | 1.85 | 1.87 | 1.835 | 1.88 | 1.9 | 1.8538 | 1.8333 | 1.855 | 2.265 | 1.9457 | 1.84 | 1.83 | 1.8681 | 1.93 | 1.845 | 1.98 | 1.5 | 1.845 | 1.8906 | 1.9412 | 1.855 | 1.655 | 1.813
| Std Dev. |  | 0.019408 | 0.054283 | 0.021602 | 0.01169 | 0.035024 | 0.013663 | 0.039497 | 0.045166 | 0.028007 | 0.016502 | 0.019408 | 0.20461 | 0.023402 | 0.025033 | 0.028048 | 0.017799 | 0.05465 | 0.024833 | 0.042895 | 0.038687 | 0.030111 | 0.36581 | 0.020032 | 0.030768 | 0.041191 | 0.01164
| Rel.Std.Dev. |  | 1.044% | 2.793% | 1.254% | 0.6325% | 1.87% | 0.7439% | 2.09% | 2.377% | 1.523% | 0.8992% | 1.048% | 8.871% | 1.204% | 1.356% | 1.53% | 0.9542% | 2.837% | 1.351% | 2.166% | 2.582% | 1.631% | 18.09% | 1.034% | 1.66% | 2.499% | 0.6433%
| PDM3 |  | -0.5603% | 3.988% | -7.784% | -1.095% | 0.2423% | -1.72% | 1.134% | 1.669% | -1.575% | -1.804% | -0.9171% | 23.43% | 4.017% | -1.185% | -1.898% | -0.1919% | 3.096% | -1.631% | 5.95% | -19.82% | -1.185% | 8.203% | 3.674% | -0.8279% | -11.8% | -3.173%
| Z-Score (Absolute) |  | 0.13 | 2.1 | 2.99 | 0.1 | 0.48 | 0.37 | 0.87 | 1.1 | 0.3 | 0.4 | 0.02 | 10.5 | 2.11 | 0.13 | 0.44 | 0.29 | 1.71 | 0.33 | 2.95 | 8.19 | 0.13 | 3.92 | 1.96 | 0.02 | 4.72 | 0.99
| Analytical results for Cd in OREAS 505b (Certified Value 0.29 ppm) | [28 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 0.31 | 0.5 | 0.25 | 0.3 | 0.29 | 0.3 | 0.28 | 0.27 | 0.28 | 0.1555 | 0.30538 | 0.34 | < 0.1 | < 1 | 0.39 | 0.29098 | 0.31 | 0.31 | 0.26 | < 0.5 | 0.25 | < 5 | 0.28 | 0.3 | 0.3 | < 1 |  | 1
| 1 | 2 | 0.3 | 0.5 | 0.26 | 0.27 | 0.29 | 0.32 | 0.28 | 0.26 | 0.28 | 0.1449 | 0.30066 | 0.34 | < 0.1 | < 1 | 0.36 | 0.31919 | 0.3 | 0.31 | 0.28 | < 0.5 | 0.3 | < 5 | 0.274 | 0.3 | 0.31 | < 1 |  | 30
| 1 | 3 | 0.29 | 0.4 | 0.25 | 0.3 | 0.32 | 0.31 | 0.29 | 0.3 | 0.3 | < 0.1 | 0.30846 | 0.32 | 0.12 | < 1 | 0.37 | 0.30001 | 0.29 | 0.29 | 0.26 | < 0.5 | 0.26 | < 5 | 0.281 | 0.4 | 0.28 | < 1 |  | 16
| 1 | 4 | 0.3 | 0.4 | 0.26 | 0.31 | 0.3 | 0.3 | 0.28 | 0.27 | 0.26 | 0.10293 | 0.30698 | 0.35 | < 0.1 | < 1 | 0.37 | 0.30386 | 0.31 | 0.27 | 0.27 | < 0.5 | 0.31 | < 5 | 0.262 | 0.3 | 0.33 | < 1 |  | 0.29222
| 1 | 5 | 0.31 | 0.5 | 0.27 | 0.29 | 0.29 | 0.33 | 0.28 | 0.25 | 0.33 | 0.10951 | 0.3179 | 0.3 | < 0.1 | < 1 | 0.35 | 0.28492 | 0.29 | 0.32 | 0.28 | < 0.5 | 0.24 | < 5 | 0.283 | 0.4 | 0.32 | < 1 |  | 18
| 1 | 6 | 0.29 | 0.5 | 0.26 | 0.29 | 0.31 | 0.26 | 0.28 | 0.3 | 0.29 | 0.10093 | 0.28449 | 0.3 | 0.11 | < 1 | 0.37 | 0.32473 | 0.3 | 0.3 | 0.27 | < 0.5 | 0.28 | < 5 | 0.275 | 0.3 | 0.32 | < 1
| Mean |  | 0.3 | 0.46667 | 0.25833 | 0.29333 | 0.3 | 0.30333 | 0.28167 | 0.275 | 0.29 | 0.12275 | 0.30398 | 0.325 | 0.115 |  | 0.36833 | 0.30395 | 0.3 | 0.3 | 0.27 |  | 0.27333 |  | 0.27583 | 0.33333 | 0.31
| Median |  | 0.3 | 0.5 | 0.26 | 0.295 | 0.295 | 0.305 | 0.28 | 0.27 | 0.285 | 0.10951 | 0.30618 | 0.33 | 0.115 |  | 0.37 | 0.30193 | 0.3 | 0.305 | 0.27 |  | 0.27 |  | 0.2775 | 0.3 | 0.315
| Std Dev. |  | 0.0089443 | 0.05164 | 0.0075277 | 0.013663 | 0.012649 | 0.024221 | 0.0040825 | 0.020736 | 0.023664 | 0.025532 | 0.011097 | 0.021679 | 0.0070711 |  | 0.013292 | 0.015556 | 0.0089443 | 0.017889 | 0.0089443 |  | 0.028048 |  | 0.0076267 | 0.05164 | 0.017889
| Rel.Std.Dev. |  | 2.981% | 11.07% | 2.914% | 4.658% | 4.216% | 7.985% | 1.449% | 7.541% | 8.16% | 20.8% | 3.651% | 6.671% | 6.149% |  | 3.609% | 5.118% | 2.981% | 5.963% | 3.313% |  | 10.26% |  | 2.765% | 15.49% | 5.77%
| PDM3 |  | 2.664% | 59.7% | -11.59% | 0.3827% | 2.664% | 3.805% | -3.61% | -5.891% | -0.758% | -57.99% | 4.025% | 11.22% | -60.65% |  | 26.05% | 4.015% | 2.664% | 2.664% | -7.602% |  | -6.462% |  | -5.606% | 14.07% | 6.086%
| Z-Score (Absolute) |  | 0 | NA | 1.14 | 0.18 | 0 | 0.09 | 0.5 | 0.69 | 0.27 | 5.19 | 0.11 | 0.69 | 6.26 | 5.49 | 1.87 | 0.11 | 0 | 0 | 0.82 | 1.37 | 0.73 | 60.34 | 0.66 | NA | 0.27 | 5.49
Results from laboratories 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 505b (Certified Value 68 ppm) | [28 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 0
1
| 1 | 1 | 79 | 62.08 | 70.39 | 70.8 | 66.9 | 72.4 | 68.2 | 71.64 | 56.0865 | 69.9179 | 83.6 | 69.1 | 64 | 71.9 | 69.4 | 67.9 | 56.866 | 40.041 | 74 | 60.53 | [7 empty] | 1
| 1 | 2 | 77 | 62.49 | 75.81 | 69.2 | 67.5 | 70.4 | 69.6 | 69.72 | 62.4757 | 69.1836 | 81.8 | 64.4 | 64 | 69 | 71.2 | 74.11 | 58.359 | 39.634 | 71 | 54.17 | [7 empty] | 31
| 1 | 3 | 73 | 63.4 | 75.36 | 69.5 | 68.7 | 74.4 | 73.7 | 70.95 | 58.2127 | 73.7621 | 85.1 | 60.9 | 64 | 65.4 | 69.2 | 68.19 | 54.399 | 39.903 | 78 | 56.62 | [7 empty] | 16
| 1 | 4 | 75 | 61.4 | 74.67 | 71 | 69.2 | 75.4 | 68.3 | 72.68 | 58.3126 | 71.9609 | 83.1 | 67.7 | 65 | 67.3 | 72.5 | 73.72 | 56.084 | 39.637 | 75 | 59.16 | [7 empty] | 67.981
| 1 | 5 | 72 | 64.27 | 74.2 | 72.1 | 69.6 | 69.2 | 66.2 | 73.44 | 61.3689 | 71.051 | 84 | 66.3 | 67 | 71.3 | 72.3 | 65.48 | 55.984 | 39.742 | 77 | 56.17 | [7 empty] | 19
| 1 | 6 | 75 | 63.59 | 75.83 | 74.4 | 65 | 71.9 | 68.7 | 69.17 | 59.5522 | 70.3644 | 85.9 | 72.6 | 65 | 65.6 | 69.9 | 68.29 | 56.992 | 40.083 | 75 | 58
| Mean |  | 75.1667 | 62.8717 | 74.3767 | 71.1667 | 67.8167 | 72.2833 | 69.1167 | 71.2667 | 59.3348 | 71.04 | 83.9167 | 66.8333 | 64.8333 | 68.4167 | 70.75 | 69.615 | 56.4473 | 39.84 | 75 | 57.4417
| Median |  | 75 | 62.945 | 75.015 | 70.9 | 68.1 | 72.15 | 68.5 | 71.295 | 58.9324 | 70.7077 | 83.8 | 67 | 64.5 | 68.15 | 70.55 | 68.24 | 56.475 | 39.8225 | 75 | 57.31
| Std Dev. |  | 2.5626 | 1.0666 | 2.0564 | 1.9044 | 1.7175 | 2.3413 | 2.5071 | 1.6582 | 2.3197 | 1.6383 | 1.4552 | 4.0138 | 1.169 | 2.7953 | 1.4571 | 3.4895 | 1.3167 | 0.19833 | 2.4495 | 2.2715
| Rel.Std.Dev. |  | 3.409% | 1.697% | 2.765% | 2.676% | 2.533% | 3.239% | 3.627% | 2.327% | 3.91% | 2.306% | 1.734% | 6.006% | 1.803% | 4.086% | 2.059% | 5.013% | 2.333% | 0.4978% | 3.266% | 3.954%
| PDM3 |  | 10.57% | -7.516% | 9.408% | 4.686% | -0.2417% | 6.329% | 1.671% | 4.833% | -12.72% | 4.5% | 23.44% | -1.688% | -4.63% | 0.6409% | 4.073% | 2.404% | -16.97% | -41.4% | 10.32% | -15.5%
| Z-Score (Absolute) |  | 1.05 | 1.18 | 0.91 | 0.33 | 0.28 | 0.53 | 0.05 | 0.34 | 1.82 | 0.3 | 2.63 | 0.46 | 0.82 | 0.17 | 0.25 | 0.05 | 2.34 | 5.34 | 1.02 | 2.16
| Analytical results for Co in OREAS 505b (Certified Value 9.18 ppm) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 2
3
| 1 | 1 | 9 | 10.6 | 9.2 | 9.3 | 9 | 8.9 | 9.3 | 8.8 | 9.4 | 8.0539 | 9.6075 | 9.63 | 10.19 | 12.63 | 10.41 | 9.9 | 9.4113 | 9.2 | 8.7 | 9.8 | 12.2 | 8.7 | 9.211 | 8.4635 | 9.2 | 9.2 | 7 | 1
| 1 | 2 | 9 | 10.1 | 9.3 | 9.5 | 9.1 | 9.2 | 9.6 | 8.8 | 9.2 | 8.0687 | 9.1517 | 9.89 | 8.92 | 13.44 | 9.99 | 9.1 | 9.6307 | 8.9 | 7.4 | 8.5 | 12.3 | 9 | 8.987 | 8.0897 | 9.1 | 8.9 | 7 | 17
| 1 | 3 | 9 | 10.2 | 9.2 | 9.5 | 9.2 | 8.9 | 10.1 | 9.3 | 9.1 | 8.2053 | 9.1847 | 9.21 | 10.45 | 13.77 | 10 | 9.3 | 9.4716 | 8.4 | 7.9 | 8.2 | 12.2 | 9.5 | 9.194 | 8.2898 | 9.5 | 8.7 | 8 | 16
| 1 | 4 | 9 | 10.1 | 9.1 | 9.6 | 9.4 | 9 | 9.9 | 8.7 | 9.5 | 8.2833 | 9.2251 | 10 | 9.87 | 14.29 | 10.24 | 9.5 | 9.6492 | 8.7 | 9.2 | 8.6 | 12.3 | 8.4 | 9.431 | 8.2878 | 9 | 9.4 | 7 | 9.1779
| 1 | 5 | 10 | 10.3 | 9.3 | 9.2 | 9.5 | 9.2 | 9.3 | 8.7 | 9.6 | 8.0208 | 9.0835 | 9.81 | 9.38 | 14.02 | 9.88 | 8.8 | 9.3721 | 8.6 | 8.4 | 8.5 | 12 | 8.6 | 9.51 | 8.1544 | 9.2 | 9.2 | 7 | 20
| 1 | 6 | 10 | 10.4 | 9.3 | 9.2 | 9.2 | 8.5 | 9.8 | 9.3 | 9.8 | 8.2168 | 9.1775 | 10.3 | 10.45 | 14.23 | 10.03 | 9.7 | 9.4769 | 8.5 | 8.9 | 8.2 | 12.4 | 8.6 | 9.354 | 8.2135 | 9.4 | 8.9 | 7
| Mean |  | 9.3333 | 10.2833 | 9.2333 | 9.3833 | 9.2333 | 8.95 | 9.6667 | 8.9333 | 9.4333 | 8.1415 | 9.2383 | 9.8067 | 9.8767 | 13.73 | 10.0917 | 9.3833 | 9.502 | 8.7167 | 8.4167 | 8.6333 | 12.2333 | 8.8 | 9.2812 | 8.2498 | 9.2333 | 9.05 | 7.1667
| Median |  | 9 | 10.25 | 9.25 | 9.4 | 9.2 | 8.95 | 9.7 | 8.8 | 9.45 | 8.137 | 9.1811 | 9.85 | 10.03 | 13.895 | 10.015 | 9.4 | 9.4743 | 8.65 | 8.55 | 8.5 | 12.25 | 8.65 | 9.2825 | 8.2507 | 9.2 | 9.05 | 7
| Std Dev. |  | 0.5164 | 0.19408 | 0.08165 | 0.17224 | 0.18619 | 0.25884 | 0.3266 | 0.28752 | 0.2582 | 0.10714 | 0.18682 | 0.36741 | 0.61928 | 0.62344 | 0.19529 | 0.40208 | 0.11392 | 0.29269 | 0.66758 | 0.59554 | 0.13663 | 0.39497 | 0.1893 | 0.13021 | 0.18619 | 0.25884 | 0.40825
| Rel.Std.Dev. |  | 5.533% | 1.887% | 0.8843% | 1.836% | 2.016% | 2.892% | 3.379% | 3.218% | 2.737% | 1.316% | 2.022% | 3.746% | 6.27% | 4.541% | 1.935% | 4.285% | 1.199% | 3.358% | 7.932% | 6.898% | 1.117% | 4.488% | 2.04% | 1.578% | 2.016% | 2.86% | 5.696%
| PDM3 |  | 1.694% | 12.05% | 0.6045% | 2.239% | 0.6045% | -2.483% | 5.326% | -2.664% | 2.784% | -11.29% | 0.6588% | 6.851% | 7.614% | 49.6% | 9.957% | 2.239% | 3.532% | -5.025% | -8.294% | -5.933% | 33.29% | -4.117% | 1.126% | -10.11% | 0.6045% | -1.393% | -21.91%
| Z-Score (Absolute) |  | NA | 1.65 | 0.01 | 0.23 | 0.01 | 0.45 | 0.67 | 0.48 | 0.31 | 1.73 | 0 | 0.89 | 1 | 7.07 | 1.34 | 0.23 | 0.42 | 0.82 | 1.29 | 0.95 | 4.71 | 0.69 | 0.07 | 1.56 | 0.01 | 0.3 | NA
Results from laboratories 1.01 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 505b (Certified Value 58 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 60 | 51 | 55 | 40 | 58 | 62 | 59 | 59 | 64 | 59.4084 | 30 | 50.18 | 67.63 | 65.08 | 48 | 63.4337 | 48 | 59 | 57 | 56.6 | 63 | 59.904 | 56.9773 | 49 | 57 | 57 |  | 1
| 1 | 2 | 60 | 52 | 52 | 42 | 61 | 61 | 58 | 60 | 54 | 61.1125 | 32.1 | 46.06 | 72.43 | 66.04 | 48 | 60.0321 | 56 | 59 | 58 | 56.7 | 59 | 60.853 | 55.7805 | 50 | 59 | 49 |  | 18
| 1 | 3 | 58 | 51 | 52 | 41 | 60 | 61 | 56 | 56 | 66 | 58.7832 | 35.1 | 55.02 | 78.22 | 67.71 | 48 | 61.7599 | 69 | 57 | 59 | 57.2 | 57 | 62.239 | 56.6768 | 53 | 57 | 51 |  | 16
| 1 | 4 | 62 | 52 | 55 | 40 | 60 | 61 | 61 | 60 | 56 | 58.7033 | 32.4 | 50.92 | 77.29 | 67.78 | 46 | 58.8593 | 67 | 59 | 60 | 59.3 | 54 | 59.866 | 57.0789 | 52 | 59 | 46 |  | 58.1552
| 1 | 5 | 61 | 52 | 53 | 38 | 61 | 61 | 61 | 58 | 68 | 58.6284 | 31 | 57.82 | 74.86 | 65.42 | 47 | 58.5296 | 62 | 59 | 61 | 58.1 | 54 | 61.769 | 55.8797 | 51 | 59 | 45 |  | 21
| 1 | 6 | 63 | 54 | 54 | 43 | 59 | 61 | 62 | 58 | 68 | 60.496 | 30.2 | 55.38 | 64.47 | 65.92 | 50 | 62.1107 | 71 | 57 | 60 | 58.8 | 57 | 62.102 | 56.4795 | 54 | 57 | 55
| Mean |  | 60.6667 | 52 | 53.5 | 40.6667 | 59.8333 | 61.1667 | 59.5 | 58.5 | 62.6667 | 59.522 | 31.8 | 52.5633 | 72.4833 | 66.325 | 47.8333 | 60.7875 | 62.1667 | 58.3333 | 59.1667 | 57.7833 | 57.3333 | 61.1222 | 56.4788 | 51.5 | 58 | 50.5
| Median |  | 60.5 | 52 | 53.5 | 40.5 | 60 | 61 | 60 | 58.5 | 65 | 59.0958 | 31.55 | 52.97 | 73.645 | 65.98 | 48 | 60.896 | 64.5 | 59 | 59.5 | 57.65 | 57 | 61.311 | 56.5782 | 51.5 | 58 | 50
| Std Dev. |  | 1.7512 | 1.0954 | 1.3784 | 1.7512 | 1.169 | 0.40825 | 2.2583 | 1.5166 | 6.1536 | 1.0494 | 1.8857 | 4.2959 | 5.4669 | 1.1532 | 1.3292 | 1.9538 | 8.7958 | 1.0328 | 1.472 | 1.1268 | 3.3862 | 1.0733 | 0.54669 | 1.8708 | 1.0954 | 4.8062
| Rel.Std.Dev. |  | 2.887% | 2.107% | 2.576% | 4.306% | 1.954% | 0.6674% | 3.795% | 2.592% | 9.82% | 1.763% | 5.93% | 8.173% | 7.542% | 1.739% | 2.779% | 3.214% | 14.15% | 1.771% | 2.488% | 1.95% | 5.906% | 1.756% | 0.968% | 3.633% | 1.889% | 9.517%
| PDM3 |  | 4.318% | -10.58% | -8.005% | -30.07% | 2.886% | 5.178% | 2.312% | 0.5928% | 7.758% | 2.35% | -45.32% | -9.615% | 24.64% | 14.05% | -17.75% | 4.526% | 6.898% | 0.3062% | 1.739% | -0.6395% | -1.413% | 5.102% | -2.883% | -11.44% | -0.267% | -13.16%
| Z-Score (Absolute) |  | 0.56 | 1.59 | 1.22 | 4.39 | 0.35 | 0.68 | 0.27 | 0.02 | 1.05 | 0.27 | 6.58 | 1.45 | 3.48 | 1.95 | 2.62 | 0.59 | 0.93 | 0.02 | 0.19 | 0.16 | 0.27 | 0.67 | 0.48 | 1.71 | 0.1 | 1.96
| Analytical results for Cs in OREAS 505b (Certified Value 9.54 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
3
| 1 | 1 | 9.9 | 8.5 | 9.59 | 9.16 | 9.34 | 10.45 | 9.17 | 9.77 | 10.2273 | 9.91 | 10 | 9.9289 | 9.43 | 10.05 | 9.97 | 9.8 | 8.463 | 7.3 | 9.12 | [8 empty] | 1
| 1 | 2 | 9.4 | 8.5 | 9.4 | 9.23 | 9.62 | 10.75 | 9.34 | 9.76 | 10.2933 | 10.1 | 9 | 9.8797 | 9.32 | 10.2 | 9.88 | 10.1 | 8.224 | 7.4 | 8.81 | [8 empty] | 34
| 1 | 3 | 9.2 | 8.8 | 9.49 | 9.32 | 9.42 | 11.05 | 9.42 | 9.81 | 10.4068 | 9.87 | 9 | 9.8874 | 9.05 | 9.11 | 9.72 | 9.8 | 8.309 | 8 | 8.83 | [8 empty] | 16
| 1 | 4 | 9.3 | 8.3 | 9.57 | 9.57 | 9.58 | 11 | 9.33 | 10.11 | 10.4963 | 10.1 | 10 | 10.0376 | 9.2 | 9.39 | 10.05 | 9.9 | 8.313 | 7.9 | 9.58 | [8 empty] | 9.5392
| 1 | 5 | 9.4 | 8.5 | 9.16 | 9.63 | 9.65 | 10.35 | 9.06 | 9.7 | 10.2343 | 9.94 | 9 | 9.9206 | 9.14 | 9.52 | 10.1 | 9.7 | 8.143 | 7.7 | 9.34 | [8 empty] | 22
| 1 | 6 | 9.2 | 8.5 | 9.26 | 9.43 | 9.37 | 10.85 | 9.35 | 9.99 | 10.2762 | 10.6 | 10 | 10.0385 | 9.09 | 9.67 | 9.95 | 9.7 | 8.478 | 8.2 | 8.41
| Mean |  | 9.4 | 8.5167 | 9.4117 | 9.39 | 9.4967 | 10.7417 | 9.2783 | 9.8567 | 10.3224 | 10.0867 | 9.5 | 9.9488 | 9.205 | 9.6567 | 9.945 | 9.8333 | 8.3217 | 7.75 | 9.015
| Median |  | 9.35 | 8.5 | 9.445 | 9.375 | 9.5 | 10.8 | 9.335 | 9.79 | 10.2847 | 10.02 | 9.5 | 9.9248 | 9.17 | 9.595 | 9.96 | 9.8 | 8.311 | 7.8 | 8.975
| Std Dev. |  | 0.26077 | 0.16021 | 0.17291 | 0.18708 | 0.13574 | 0.28708 | 0.13497 | 0.15845 | 0.10693 | 0.26964 | 0.54772 | 0.071665 | 0.14516 | 0.40957 | 0.13457 | 0.15055 | 0.13123 | 0.35071 | 0.41908
| Rel.Std.Dev. |  | 2.774% | 1.881% | 1.837% | 1.992% | 1.429% | 2.673% | 1.455% | 1.608% | 1.036% | 2.673% | 5.766% | 0.7203% | 1.577% | 4.241% | 1.353% | 1.531% | 1.577% | 4.525% | 4.649%
| PDM3 |  | -1.459% | -10.72% | -1.337% | -1.564% | -0.4461% | 12.61% | -2.735% | 3.328% | 8.21% | 5.739% | -0.4112% | 4.294% | -3.504% | 1.231% | 4.254% | 3.083% | -12.76% | -18.76% | -5.495%
| Z-Score (Absolute) |  | 0.09 | 1.5 | 0.07 | 0.1 | 0.07 | 2.06 | 0.28 | 0.64 | 1.39 | 1.01 | NA | 0.79 | 0.4 | 0.32 | 0.79 | 0.61 | 1.81 | 2.73 | 0.7
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 505b (Certified Value 0.32 wt.%) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 3
3
| 1 | 1 | 0.317 | 0.32882 | 0.30173 | 0.3317 | 0.299 | 0.317 | 0.322 | 0.32 | 0.32314 | 0.3064 | 0.31295 | 0.328 | 0.31262 | 0.31895 | 0.3254 | 0.3276 | 0.32677 | 0.3292 | 0.314 | 0.325 | 0.30417 | 0.32028 | 0.33952 | 0.31883 | 0.3122 | 0.3293 | 0.3097 | 1
| 1 | 2 | 0.317 | 0.31775 | 0.30924 | 0.3232 | 0.311 | 0.327 | 0.328 | 0.317 | 0.32373 | 0.31188 | 0.31315 | 0.328 | 0.31451 | 0.31039 | 0.3217 | 0.3202 | 0.32951 | 0.3211 | 0.323 | 0.334 | 0.30864 | 0.32334 | 0.33722 | 0.31094 | 0.31248 | 0.3271 | 0.3126 | 19
| 1 | 3 | 0.314 | 0.32597 | 0.30874 | 0.3388 | 0.315 | 0.318 | 0.325 | 0.318 | 0.31907 | 0.30541 | 0.30757 | 0.321 | 0.32524 | 0.31057 | 0.3204 | 0.3241 | 0.32141 | 0.3216 | 0.318 | 0.324 | 0.30239 | 0.32028 | 0.32911 | 0.31957 | 0.31748 | 0.3248 | 0.3108 | 16
| 1 | 4 | 0.319 | 0.3319 | 0.30579 | 0.3174 | 0.304 | 0.324 | 0.335 | 0.324 | 0.33125 | 0.30324 | 0.31126 | 0.326 | 0.32533 | 0.31513 | 0.3204 | 0.324 | 0.32595 | 0.3226 | 0.319 | 0.338 | 0.30891 | 0.31722 | 0.31957 | 0.31892 | 0.31279 | 0.3242 | 0.3102 | 0.31974
| 1 | 5 | 0.319 | 0.32382 | 0.30645 | 0.3215 | 0.326 | 0.321 | 0.317 | 0.318 | 0.33688 | 0.30993 | 0.30891 | 0.319 | 0.33415 | 0.31924 | 0.325 | 0.3248 | 0.32267 | 0.3257 | 0.319 | 0.338 | 0.30604 | 0.3264 | 0.34145 | 0.31485 | 0.31234 | 0.3215 | 0.3083 | 23
| 1 | 6 | 0.318 | 0.33731 | 0.30429 | 0.3195 | 0.312 | 0.313 | 0.33 | 0.314 | 0.32388 | 0.30734 | 0.31314 | 0.329 | 0.32523 | 0.30999 | 0.327 | 0.3275 | 0.33094 | 0.3234 | 0.317 | 0.322 | 0.30603 | 0.32232 | 0.34276 | 0.31773 | 0.31913 | 0.3276 | 0.307
| Mean |  | 0.31733 | 0.3276 | 0.30604 | 0.32535 | 0.31117 | 0.32 | 0.32617 | 0.3185 | 0.32632 | 0.30737 | 0.31116 | 0.32517 | 0.32285 | 0.31405 | 0.32332 | 0.3247 | 0.32621 | 0.32393 | 0.31833 | 0.33017 | 0.30603 | 0.32164 | 0.33494 | 0.31681 | 0.3144 | 0.32575 | 0.30977
| Median |  | 0.3175 | 0.32739 | 0.30612 | 0.32235 | 0.3115 | 0.3195 | 0.3265 | 0.318 | 0.32381 | 0.30687 | 0.3121 | 0.327 | 0.32523 | 0.31285 | 0.32335 | 0.32445 | 0.32636 | 0.323 | 0.3185 | 0.3295 | 0.30603 | 0.3213 | 0.33837 | 0.31828 | 0.31263 | 0.32595 | 0.30995
| Std Dev. |  | 0.0018619 | 0.0067562 | 0.0028071 | 0.0082231 | 0.0093256 | 0.0050596 | 0.0063061 | 0.0033317 | 0.0064985 | 0.0031224 | 0.0024106 | 0.0041673 | 0.0079914 | 0.0043382 | 0.0028414 | 0.0027335 | 0.0037194 | 0.0030461 | 0.0029439 | 0.0073326 | 0.0025241 | 0.0031383 | 0.0089421 | 0.0033265 | 0.0030732 | 0.0028005 | 0.0019541
| Rel.Std.Dev. |  | 0.5867% | 2.062% | 0.9172% | 2.527% | 2.997% | 1.581% | 1.933% | 1.046% | 1.991% | 1.016% | 0.7747% | 1.282% | 2.475% | 1.381% | 0.8788% | 0.8419% | 1.14% | 0.9403% | 0.9248% | 2.221% | 0.8248% | 0.9757% | 2.67% | 1.05% | 0.9775% | 0.8597% | 0.6308%
| PDM3 |  | -0.7523% | 2.457% | -4.284% | 1.755% | -2.681% | 0.08174% | 2.01% | -0.3874% | 2.06% | -3.869% | -2.683% | 1.698% | 0.9719% | -1.781% | 1.119% | 1.552% | 2.023% | 1.312% | -0.4395% | 3.261% | -4.288% | 0.5947% | 4.754% | -0.9168% | -1.669% | 1.88% | -3.119%
| Z-Score (Absolute) |  | 0.64 | 0.88 | 2.3 | 0.55 | 1.55 | 0.24 | 0.67 | 0.46 | 0.69 | 2.11 | 1.55 | 0.52 | 0.18 | 1.12 | 0.25 | 0.45 | 0.67 | 0.34 | 0.49 | 1.26 | 2.3 | 0 | 1.96 | 0.71 | 1.07 | 0.61 | 1.75
| Analytical results for Dy in OREAS 505b (Certified Value 3.3 ppm) | [28 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS | 2
2
| 1 | 1 | 1.01: 3.3; 1.02: 3.46; 1.03: 2.5828; 1.04: 3.7334; 1.05: 3.62; 1.06: 3.4; 1.07: 3.027 | 1
| 1 | 2 | 1.01: 3.3; 1.02: 3.45; 1.03: 2.9876; 1.04: 3.7237; 1.05: 3.53; 1.06: 3.7; 1.07: 2.985 | 6
| 1 | 3 | 1.01: 3.4; 1.02: 3.53; 1.03: 2.6898; 1.04: 3.7153; 1.05: 3.37; 1.06: 3.5; 1.07: 3.054 | 16
| 1 | 4 | 1.01: 3.2; 1.02: 3.4; 1.03: 2.6897; 1.04: 3.7104; 1.05: 3.48; 1.06: 3.5; 1.07: 3.04 | 3.3024
| 1 | 5 | 1.01: 3.3; 1.02: 3.39; 1.03: 2.878; 1.04: 3.7775; 1.05: 3.53; 1.06: 3.5; 1.07: 2.952 | 24
| 1 | 6 | 1.01: 3.3; 1.02: 3.58; 1.03: 2.8618; 1.04: 3.7948; 1.05: 3.45; 1.06: 3.5; 1.07: 3.078
| Mean |  | 1.01: 3.3; 1.02: 3.4683; 1.03: 2.7816; 1.04: 3.7425; 1.05: 3.4967; 1.06: 3.5167; 1.07: 3.0227
| Median |  | 1.01: 3.3; 1.02: 3.455; 1.03: 2.7758; 1.04: 3.7286; 1.05: 3.505; 1.06: 3.5; 1.07: 3.0335
| Std Dev. |  | 1.01: 0.063246; 1.02: 0.07414; 1.03: 0.15136; 1.04: 0.035137; 1.05: 0.084774; 1.06: 0.098319; 1.07: 0.046423
| Rel.Std.Dev. |  | 1.01: 1.917%; 1.02: 2.138%; 1.03: 5.442%; 1.04: 0.9388%; 1.05: 2.424%; 1.06: 2.796%; 1.07: 1.536%
| PDM3 |  | 1.01: -0.07154%; 1.02: 5.026%; 1.03: -15.77%; 1.04: 13.33%; 1.05: 5.884%; 1.06: 6.489%; 1.07: -8.47%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0; 1.03: 1.69; 1.04: 0.67; 1.05: 0.07; 1.06: NA; 1.07: 1.1
| Analytical results for Er in OREAS 505b (Certified Value 1.36 ppm) | [28 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS | 2
2
| 1 | 1 | 1.01: 1.2; 1.02: 1.42; 1.03: 1.0814; 1.04: 1.5054; 1.05: 1.49; 1.06: 1.5; 1.07: 1.266 | 1
| 1 | 2 | 1.01: 1.2; 1.02: 1.43; 1.03: 1.2285; 1.04: 1.5033; 1.05: 1.5; 1.06: 1.5; 1.07: 1.265 | 7
| 1 | 3 | 1.01: 1.3; 1.02: 1.44; 1.03: 1.1401; 1.04: 1.5654; 1.05: 1.34; 1.06: 1.5; 1.07: 1.282 | 16
| 1 | 4 | 1.01: 1.2; 1.02: 1.37; 1.03: 1.0692; 1.04: 1.5238; 1.05: 1.42; 1.06: 1.5; 1.07: 1.341 | 1.3622
| 1 | 5 | 1.01: 1.2; 1.02: 1.43; 1.03: 1.1567; 1.04: 1.5495; 1.05: 1.46; 1.06: 1.4; 1.07: 1.251 | 25
| 1 | 6 | 1.01: 1.3; 1.02: 1.47; 1.03: 1.1761; 1.04: 1.51; 1.05: 1.38; 1.06: 1.5; 1.07: 1.302
| Mean |  | 1.01: 1.2333; 1.02: 1.4267; 1.03: 1.142; 1.04: 1.5262; 1.05: 1.4317; 1.06: 1.4833; 1.07: 1.2845
| Median |  | 1.01: 1.2; 1.02: 1.43; 1.03: 1.1484; 1.04: 1.5169; 1.05: 1.44; 1.06: 1.5; 1.07: 1.274
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.03266; 1.03: 0.059728; 1.04: 0.02571; 1.05: 0.063377; 1.06: 0.040825; 1.07: 0.032709
| Rel.Std.Dev. |  | 1.01: 4.187%; 1.02: 2.289%; 1.03: 5.23%; 1.04: 1.685%; 1.05: 4.427%; 1.06: 2.752%; 1.07: 2.546%
| PDM3 |  | 1.01: -9.461%; 1.02: 4.731%; 1.03: -16.17%; 1.04: 12.04%; 1.05: 5.098%; 1.06: 8.891%; 1.07: -5.705%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0; 1.03: 1.93; 1.04: 0.67; 1.05: 0.03; 1.06: NA; 1.07: 0.96
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Eu in OREAS 505b (Certified Value 1.36 ppm) | [28 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS | 2
3
| 1 | 1 | 1.01: 1.2; 1.02: 1.36; 1.03: 1.3516; 1.04: 1.4251; 1.05: 1.3; 1.06: 1.1; 1.07: 1.271 | 1
| 1 | 2 | 1.01: 1.1; 1.02: 1.38; 1.03: 1.5536; 1.04: 1.4212; 1.05: 1.34; 1.06: 1.1; 1.07: 1.247 | 8
| 1 | 3 | 1.01: 1.2; 1.02: 1.38; 1.03: 1.4166; 1.04: 1.4285; 1.05: 1.27; 1.06: 1; 1.07: 1.285 | 16
| 1 | 4 | 1.01: 1.1; 1.02: 1.35; 1.03: 1.3907; 1.04: 1.4138; 1.05: 1.32; 1.06: 1.1; 1.07: 1.281 | 1.3608
| 1 | 5 | 1.01: 1.1; 1.02: 1.35; 1.03: 1.4903; 1.04: 1.4008; 1.05: 1.3; 1.06: 1; 1.07: 1.227 | 26
| 1 | 6 | 1.01: 1.1; 1.02: 1.42; 1.03: 1.4554; 1.04: 1.4107; 1.05: 1.29; 1.06: 1.1; 1.07: 1.294
| Mean |  | 1.01: 1.1333; 1.02: 1.3733; 1.03: 1.443; 1.04: 1.4167; 1.05: 1.3033; 1.06: 1.0667; 1.07: 1.2675
| Median |  | 1.01: 1.1; 1.02: 1.37; 1.03: 1.436; 1.04: 1.4175; 1.05: 1.3; 1.06: 1.1; 1.07: 1.276
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.026583; 1.03: 0.072677; 1.04: 0.010286; 1.05: 0.024221; 1.06: 0.05164; 1.07: 0.025548
| Rel.Std.Dev. |  | 1.01: 4.556%; 1.02: 1.936%; 1.03: 5.036%; 1.04: 0.726%; 1.05: 1.858%; 1.06: 4.841%; 1.07: 2.016%
| PDM3 |  | 1.01: -16.71%; 1.02: 0.9228%; 1.03: 6.045%; 1.04: 4.108%; 1.05: -4.221%; 1.06: -21.61%; 1.07: -6.855%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0; 1.03: 0.67; 1.04: 0.42; 1.05: 0.68; 1.06: NA; 1.07: 1.02
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Fe in OREAS 505b (Certified Value 3.04 wt.%) | [28 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 3.01 | 3.02 | 2.9 | 3.13 | 2.92 | 2.95 | 2.98 | 3.02 | 3.07 | 3.0969 | 2.89 | 3.22 | 3.3469 | 3.22 | 3.1 | 3.0679 | 3.06 | 2.94 | 2.91 | 3.14 | 3.13 | 3.1408 | 3.1238 | 3.13 | 2.83 | 2.96 |  | 1
| 1 | 2 | 3.09 | 2.89 | 2.94 | 3.11 | 3.02 | 3.05 | 3.06 | 3.07 | 3.11 | 3.0457 | 2.98 | 3 | 3.365 | 3.19 | 3.03 | 3.0479 | 3.03 | 2.97 | 3.03 | 3.17 | 3.19 | 3.0947 | 3.0484 | 3.13 | 2.85 | 2.98 |  | 12
| 1 | 3 | 3.01 | 2.98 | 2.94 | 3.09 | 3 | 2.98 | 3.06 | 3.01 | 3.18 | 2.9309 | 2.95 | 3.44 | 3.3474 | 3.21 | 3.01 | 3.0241 | 3.05 | 2.94 | 3.01 | 3.17 | 3.01 | 3.0703 | 3.1263 | 3.18 | 2.81 | 2.98 |  | 16
| 1 | 4 | 3.08 | 3.08 | 2.9 | 3.06 | 3.03 | 3.02 | 3.21 | 3.08 | 3.16 | 2.9677 | 2.98 | 3.41 | 3.3066 | 3.16 | 3.08 | 3.0192 | 3.06 | 2.96 | 3.08 | 3.19 | 3.06 | 3.1284 | 3.127 | 3.14 | 2.85 | 2.98 |  | 3.0429
| 1 | 5 | 3.1 | 2.95 | 2.98 | 3.09 | 3.08 | 2.98 | 3.01 | 3 | 3.25 | 2.9731 | 2.9 | 3.37 | 3.3174 | 3.19 | 3.03 | 3.0614 | 3.09 | 2.96 | 3.11 | 3.14 | 3.09 | 3.1318 | 3.062 | 3.13 | 2.86 | 2.93 |  | 27
| 1 | 6 | 3.03 | 3.06 | 2.93 | 3.09 | 3.01 | 2.97 | 3.1 | 2.98 | 3.21 | 2.9779 | 2.96 | 3.36 | 3.3567 | 3.18 | 3.08 | 3.0679 | 3.08 | 2.87 | 2.95 | 3.24 | 3.07 | 3.1249 | 3.0939 | 3.19 | 2.88 | 2.95
| Mean |  | 3.0533 | 2.9967 | 2.9317 | 3.095 | 3.01 | 2.9917 | 3.07 | 3.0267 | 3.1633 | 2.9987 | 2.9433 | 3.3 | 3.34 | 3.1917 | 3.055 | 3.0481 | 3.0617 | 2.94 | 3.015 | 3.175 | 3.0917 | 3.1152 | 3.0969 | 3.15 | 2.8467 | 2.9633
| Median |  | 3.055 | 3 | 2.935 | 3.09 | 3.015 | 2.98 | 3.06 | 3.015 | 3.17 | 2.9755 | 2.955 | 3.365 | 3.3472 | 3.19 | 3.055 | 3.0546 | 3.06 | 2.95 | 3.02 | 3.17 | 3.08 | 3.1267 | 3.1089 | 3.135 | 2.85 | 2.97
| Std Dev. |  | 0.041312 | 0.071181 | 0.029944 | 0.023452 | 0.052154 | 0.03656 | 0.080498 | 0.039833 | 0.065625 | 0.06083 | 0.039328 | 0.16529 | 0.022946 | 0.02137 | 0.036194 | 0.021808 | 0.02137 | 0.036332 | 0.075829 | 0.037283 | 0.062102 | 0.026971 | 0.034894 | 0.027568 | 0.024221 | 0.020656
| Rel.Std.Dev. |  | 1.353% | 2.375% | 1.021% | 0.7577% | 1.733% | 1.222% | 2.622% | 1.316% | 2.075% | 2.029% | 1.336% | 5.009% | 0.687% | 0.6695% | 1.185% | 0.7155% | 0.698% | 1.236% | 2.515% | 1.174% | 2.009% | 0.8658% | 1.127% | 0.8752% | 0.8509% | 0.697%
| PDM3 |  | 0.3416% | -1.521% | -3.657% | 1.711% | -1.082% | -1.685% | 0.8893% | -0.5347% | 3.957% | -1.454% | -3.273% | 8.448% | 9.762% | 4.888% | 0.3964% | 0.1687% | 0.6155% | -3.383% | -0.9181% | 4.34% | 1.601% | 2.374% | 1.774% | 3.518% | -6.45% | -2.616%
| Z-Score (Absolute) |  | 0.01 | 0.65 | 1.39 | 0.46 | 0.5 | 0.71 | 0.18 | 0.31 | 1.24 | 0.63 | 1.26 | 2.8 | 3.25 | 1.56 | 0.01 | 0.07 | 0.09 | 1.3 | 0.45 | 1.38 | 0.43 | 0.69 | 0.49 | 1.09 | 2.36 | 1.03
| Analytical results for Ga in OREAS 505b (Certified Value 19.7 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 20.5 | 17.72 | 19.37 | 18.55 | 19.3 | 19.5 | 19.1 | 20.58 | 19.1702 | 31.2 | 19.27 | 23.72 | 21 | 20.7279 | 20 | 19.5 | 18.05 | 19.6 | 20.52 | 21.682 | 15.231 | 21 | 19 |  |  |  |  | 1
| 1 | 2 | 20.5 | 17.96 | 19.14 | 18.9 | 19.3 | 20.6 | 19.15 | 20.54 | 19.5428 | 29.3 | 17.39 | 23.63 | 19.7 | 20.0975 | 20 | 20 | 17.85 | 19.7 | 20.79 | 22.102 | 14.4 | 21 | 19 |  |  |  |  | 39
| 1 | 3 | 20.6 | 17.8 | 19.74 | 19.15 | 19.7 | 21.3 | 19.95 | 20.62 | 19.428 | 31.8 | 21.65 | 24.14 | 20 | 20.7411 | 20 | 18.4 | 17.3 | 19.4 | 19.8 | 21.908 | 14.686 | 22 | 18 |  |  |  |  | 16
| 1 | 4 | 20.4 | 17.72 | 19.73 | 19.85 | 19.75 | 20.9 | 18.95 | 21.62 | 19.629 | 29.4 | 19.74 | 24.69 | 21 | 20.5752 | 19 | 18.95 | 18.3 | 19.9 | 20.52 | 21.759 | 15.575 | 21 | 20 |  |  |  |  | 19.7327
| 1 | 5 | 20.7 | 17.86 | 19.16 | 19.8 | 19.65 | 20 | 18.65 | 21.42 | 19.2723 | 29.7 | 20.62 | 23.6 | 19.8 | 20.2735 | 19 | 19 | 18.55 | 19.2 | 19.98 | 22.182 | 14.694 | 21 | 19 |  |  |  |  | 28
| 1 | 6 | 20.5 | 17.78 | 19.32 | 19.25 | 18.55 | 20.4 | 19.8 | 20.75 | 19.2688 | 30.1 | 20.2 | 24.17 | 21.1 | 21.0177 | 20 | 19.95 | 18.25 | 20.2 | 19.98 | 21.907 | 15.393 | 22 | 18
| Mean |  | 20.5333 | 17.8067 | 19.41 | 19.25 | 19.375 | 20.45 | 19.2667 | 20.9217 | 19.3852 | 30.25 | 19.8117 | 23.9917 | 20.4333 | 20.5722 | 19.6667 | 19.3 | 18.05 | 19.6667 | 20.265 | 21.9233 | 14.9965 | 21.3333 | 18.8333
| Median |  | 20.5 | 17.79 | 19.345 | 19.2 | 19.475 | 20.5 | 19.125 | 20.685 | 19.3502 | 29.9 | 19.97 | 23.93 | 20.5 | 20.6516 | 20 | 19.25 | 18.15 | 19.65 | 20.25 | 21.9075 | 14.9625 | 21 | 19
| Std Dev. |  | 0.10328 | 0.091797 | 0.26698 | 0.50695 | 0.45028 | 0.64109 | 0.50465 | 0.47305 | 0.17811 | 1.0252 | 1.4385 | 0.42414 | 0.66533 | 0.3364 | 0.5164 | 0.62849 | 0.43704 | 0.3559 | 0.39607 | 0.19214 | 0.46703 | 0.5164 | 0.75277
| Rel.Std.Dev. |  | 0.503% | 0.5155% | 1.375% | 2.634% | 2.324% | 3.135% | 2.619% | 2.261% | 0.9188% | 3.389% | 7.261% | 1.768% | 3.256% | 1.635% | 2.626% | 3.256% | 2.421% | 1.81% | 1.954% | 0.8764% | 3.114% | 2.421% | 3.997%
| PDM3 |  | 4.058% | -9.76% | -1.635% | -2.446% | -1.813% | 3.635% | -2.362% | 6.026% | -1.761% | 53.3% | 0.4003% | 21.58% | 3.551% | 4.254% | -0.3345% | -2.193% | -8.527% | -0.3345% | 2.698% | 11.1% | -24% | 8.112% | -4.558%
| Z-Score (Absolute) |  | 0.75 | 1.62 | 0.22 | 0.36 | 0.25 | 0.68 | 0.35 | 1.09 | 0.24 | 9.21 | 0.13 | 3.76 | 0.67 | 0.79 | 0 | 0.32 | 1.41 | 0 | 0.52 | 1.96 | 4.06 | NA | 0.73
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 505b (Certified Value 5.2 ppm) | [28 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  |  | 1
| No | No | 1.01: 1.03; 1.03: 1.04; 1.04: 1.10; 1.05: 1.11; 1.06: 1.19; 1.07: 1.23; 1.08: 1.26 |  |  |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 1.01: 5; 1.03: 5.38; 1.04: 4.4435; 1.05: 5.8913; 1.06: 5.58; 1.07: 5.6; 1.08: 4.31 |  |  |  |  | 1
| 1 | 2 | 1.01: 4.9; 1.03: 5.53; 1.04: 5.2308; 1.05: 5.8719; 1.06: 5.58; 1.07: 5.8; 1.08: 4.253 |  |  |  |  | 9
| 1 | 3 | 1.01: 5.1; 1.03: 5.51; 1.04: 4.7012; 1.05: 5.9433; 1.06: 5.34; 1.07: 5.6; 1.08: 4.287 |  |  |  |  | 16
| 1 | 4 | 1.01: 4.9; 1.03: 5.4; 1.04: 4.6482; 1.05: 6.1267; 1.06: 5.5; 1.07: 5.6; 1.08: 4.325 |  |  |  |  | 5.1996
| 1 | 5 | 1.01: 5; 1.03: 5.52; 1.04: 4.9527; 1.05: 5.8566; 1.06: 5.47; 1.07: 5.5; 1.08: 4.231 |  |  |  |  | 29
| 1 | 6 | 1.01: 4.9; 1.03: 5.61; 1.04: 4.85; 1.05: 5.9794; 1.06: 5.32; 1.07: 5.4; 1.08: 4.346
| Mean |  | 1.01: 4.9667; 1.03: 5.4917; 1.04: 4.8044; 1.05: 5.9448; 1.06: 5.465; 1.07: 5.5833; 1.08: 4.292
| Median |  | 1.01: 4.95; 1.03: 5.515; 1.04: 4.7756; 1.05: 5.9173; 1.06: 5.485; 1.07: 5.6; 1.08: 4.2985
| Std Dev. |  | 1.01: 0.08165; 1.03: 0.086583; 1.04: 0.27247; 1.05: 0.10024; 1.06: 0.11345; 1.07: 0.13292; 1.08: 0.043809
| Rel.Std.Dev. |  | 1.01: 1.644%; 1.03: 1.577%; 1.04: 5.671%; 1.05: 1.686%; 1.06: 2.076%; 1.07: 2.381%; 1.08: 1.021%
| PDM3 |  | 1.01: -4.48%; 1.03: 5.617%; 1.04: -7.6%; 1.05: 14.33%; 1.06: 5.105%; 1.07: 7.38%; 1.08: -17.45%
| Z-Score (Absolute) |  | 1.01: NA; 1.03: 0.04; 1.04: 0.93; 1.05: 0.67; 1.06: 0; 1.07: NA; 1.08: 1.65
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ge in OREAS 505b (Indicative Value 0.23 ppm) | [28 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.18 | 1.19 | 1.20 | 1.21 | 1.26 | 1.27 | 1.29 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [11 empty] | 2
2
| 1 | 1 | 1.9 | 0.33 | 0.24 | 0.14 | 0.2 | 0.11 | 1.8 | < 0.2 | 1.08 | < 0.1 | 0.16 | 0.14 | 0.28 | 0.079 | 4.4 | 0.4 | [11 empty] | 1
| 1 | 2 | 1.8 | 0.44 | 0.22 | 0.15 | 0.19 | 0.1 | 1.6 | < 0.2 | 0.72 | < 0.1 | 0.16 | 0.13 | 0.3 | 0.185 | 3.9 | 0.6 | [11 empty] | 4
| 1 | 3 | 1.7 | 0.38 | 0.25 | 0.14 | 0.2 | 0.11 | 1.6 | < 0.2 | 0.84 | < 0.1 | 0.18 | 0.12 | 0.25 | 0.174 | 4.4 | 0.5 | [11 empty] | 16
| 1 | 4 | 1.8 | 0.34 | 0.24 | 0.13 | 0.22 | 0.12 | 1.7 | < 0.2 | 0.54 | < 0.1 | 0.18 | 0.13 | 0.31 | 0.219 | 4.3 | 0.5 | [11 empty] | 0.22983
| 1 | 5 | 1.9 | 0.35 | 0.25 | 0.14 | 0.22 | 0.11 | 1.7 | < 0.2 | 0.81 | < 0.1 | 0.17 | 0.13 | 0.32 | 0.205 | 4.4 | 0.5 | [11 empty] | 10
| 1 | 6 | 1.8 | 0.41 | 0.25 | 0.13 | 0.2 | 0.1 | 1.8 | < 0.2 | 0.78 | < 0.1 | 0.17 | 0.15 | 0.34 | 0.138 | 4.9 | 0.4
| Mean |  | 1.8167 | 0.375 | 0.24167 | 0.13833 | 0.205 | 0.10833 | 1.7 |  | 0.795 |  | 0.17 | 0.13333 | 0.3 | 0.16667 | 4.3833 | 0.48333
| Median |  | 1.8 | 0.365 | 0.245 | 0.14 | 0.2 | 0.11 | 1.7 |  | 0.795 |  | 0.17 | 0.13 | 0.305 | 0.1795 | 4.4 | 0.5
| Std Dev. |  | 0.075277 | 0.043243 | 0.01169 | 0.0075277 | 0.012247 | 0.0075277 | 0.089443 |  | 0.1757 |  | 0.0089443 | 0.010328 | 0.031623 | 0.051196 | 0.31885 | 0.075277
| Rel.Std.Dev. |  | 4.144% | 11.53% | 4.837% | 5.442% | 5.974% | 6.949% | 5.261% |  | 22.1% |  | 5.261% | 7.746% | 10.54% | 30.72% | 7.274% | 15.57%
| PDM3 |  | 690.4% | 63.16% | 5.149% | -39.81% | -10.8% | -52.86% | 639.7% |  | 245.9% |  | -26.03% | -41.99% | 30.53% | -27.48% | 1807% | 110.3%
| Z-Score (Absolute) |  | 9.02 | 0.86 | 0.1 | 0.48 | 0.1 | 0.65 | 8.36 | 0.7 | 3.23 | 0.98 | 0.3 | 0.51 | 0.43 | 0.32 | 23.54 | 1.47
| Analytical results for Hf in OREAS 505b (Certified Value 2.04 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 2 | 1.9 | 1.77 | 2.01 | 1.9 | 2 | 2.2 | 1.9 | 2.27 | 2.1391 | 2 | 2.12 | 2.2278 | 0.9 | 2.1 | 1.9 | 2.02 | 1.923 | 2 | 2.59 | [7 empty] | 1
| 1 | 2 | 2 | 1.9 | 1.76 | 2.04 | 1.9 | 2.1 | 2.2 | 2 | 2.24 | 2.1536 | 1.9 | 1.99 | 2.2392 | 0.8 | 2 | 1.9 | 2.08 | 1.868 | 2.1 | 2.32 | [7 empty] | 25
| 1 | 3 | 2 | 1.8 | 1.83 | 2.15 | 1.9 | 2 | 2.3 | 2.1 | 2.08 | 2.2311 | 2 | 2.06 | 2.1543 | 0.8 | 2 | 1.7 | 1.93 | 1.9 | 2.1 | 2.43 | [7 empty] | 16
| 1 | 4 | 2.1 | 1.8 | 1.67 | 2.02 | 1.9 | 2 | 2.4 | 1.9 | 2.17 | 2.1508 | 1.9 | 2.04 | 2.1711 | 0.7 | 2.1 | 1.9 | 2.31 | 1.913 | 2.1 | 2.51 | [7 empty] | 2.0442
| 1 | 5 | 2.3 | 1.9 | 1.74 | 2.11 | 2 | 2.1 | 2.1 | 1.9 | 2.34 | 2.1195 | 2 | 1.95 | 2.1337 | 0.7 | 2.1 | 1.9 | 2.01 | 1.89 | 2 | 2.43 | [7 empty] | 30
| 1 | 6 | 2.2 | 1.8 | 1.74 | 2.11 | 1.8 | 2 | 2.3 | 2 | 2.2 | 2.1839 | 1.9 | 2.16 | 2.186 | 0.7 | 2 | 1.8 | 2.14 | 1.9 | 2.1 | 2.31
| Mean |  | 2.1 | 1.85 | 1.7517 | 2.0733 | 1.9 | 2.0333 | 2.25 | 1.9667 | 2.2167 | 2.163 | 1.95 | 2.0533 | 2.1853 | 0.76667 | 2.05 | 1.85 | 2.0817 | 1.899 | 2.0667 | 2.4317
| Median |  | 2.05 | 1.85 | 1.75 | 2.075 | 1.9 | 2 | 2.25 | 1.95 | 2.22 | 2.1522 | 1.95 | 2.05 | 2.1786 | 0.75 | 2.05 | 1.9 | 2.05 | 1.9 | 2.1 | 2.43
| Std Dev. |  | 0.12649 | 0.054772 | 0.051929 | 0.057504 | 0.063246 | 0.05164 | 0.10488 | 0.08165 | 0.089144 | 0.039434 | 0.054772 | 0.078401 | 0.041341 | 0.08165 | 0.054772 | 0.083666 | 0.13227 | 0.019058 | 0.05164 | 0.10815
| Rel.Std.Dev. |  | 6.023% | 2.961% | 2.965% | 2.773% | 3.329% | 2.54% | 4.661% | 4.152% | 4.022% | 1.823% | 2.809% | 3.818% | 1.892% | 10.65% | 2.672% | 4.522% | 6.354% | 1.004% | 2.499% | 4.448%
| PDM3 |  | 2.727% | -9.502% | -14.31% | 1.423% | -7.056% | -0.5338% | 10.07% | -3.795% | 8.434% | 5.808% | -4.61% | 0.4446% | 6.902% | -62.5% | 0.2815% | -9.502% | 1.831% | -7.105% | 1.097% | 18.95%
| Z-Score (Absolute) |  | 0.27 | 1.11 | 1.65 | 0.12 | 0.83 | 0.1 | 1.09 | 0.47 | 0.91 | 0.61 | 0.56 | 0.01 | 0.74 | 7.07 | 0.01 | 1.11 | 0.17 | 0.84 | 0.08 | 2.09
| Analytical results for Hg in OREAS 505b (Indicative Value 0.013 ppm) | [28 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | [7 empty] | 1
| No | No | 1.01: 1.11; 1.02: 1.19; 1.03: 1.26 | [7 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*OES/MS; 1.03: 4A*OES/MS | [7 empty] | 3
3
| 1 | 1 | 1.01: < 1; 1.02: < 2; 1.03: 0.013 | [7 empty] | 1
| 1 | 2 | 1.01: < 1; 1.02: < 2; 1.03: 0.012 | [7 empty] | 3
| 1 | 3 | 1.01: < 1; 1.02: < 2; 1.03: 0.014 | [7 empty] | 16
| 1 | 4 | 1.01: < 1; 1.02: < 2; 1.03: 0.012 | [7 empty] | 0.013
| 1 | 5 | 1.01: < 1; 1.02: < 2; 1.03: 0.013 | [7 empty] | 9
| 1 | 6 | 1.01: < 1; 1.02: < 2; 1.03: 0.014
| Mean |  | 1.03: 0.013
| Median |  | 1.03: 0.013
| Std Dev. |  | 1.03: 0.00089443
| Rel.Std.Dev. |  | 1.03: 6.88%
PDM3 | bar: 1.03
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.69; 1.03: 0.67
| Analytical results for Ho in OREAS 505b (Certified Value 0.54 ppm) | [28 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [7 empty] | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 | [7 empty] | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | 1.01: 0.5; 1.02: 0.56; 1.03: 0.3634; 1.04: 0.62503; 1.05: 0.64; 1.06: 0.6; 1.07: 0.506 | [7 empty] | 1
| 1 | 2 | 1.01: 0.5; 1.02: 0.56; 1.03: 0.4156; 1.04: 0.62142; 1.05: 0.63; 1.06: 0.6; 1.07: 0.497 | [7 empty] | 10
| 1 | 3 | 1.01: 0.5; 1.02: 0.58; 1.03: 0.3769; 1.04: 0.63457; 1.05: 0.63; 1.06: 0.6; 1.07: 0.504 | [7 empty] | 16
| 1 | 4 | 1.01: 0.5; 1.02: 0.57; 1.03: 0.3707; 1.04: 0.62953; 1.05: 0.6; 1.06: 0.6; 1.07: 0.526 | [7 empty] | 0.54333
| 1 | 5 | 1.01: 0.5; 1.02: 0.57; 1.03: 0.4027; 1.04: 0.61213; 1.05: 0.62; 1.06: 0.6; 1.07: 0.5 | [7 empty] | 31
| 1 | 6 | 1.01: 0.5; 1.02: 0.6; 1.03: 0.4017; 1.04: 0.61612; 1.05: 0.63; 1.06: 0.6; 1.07: 0.507
| Mean |  | 1.01: 0.5; 1.02: 0.57333; 1.03: 0.3885; 1.04: 0.62313; 1.05: 0.625; 1.06: 0.6; 1.07: 0.50667
| Median |  | 1.01: 0.5; 1.02: 0.57; 1.03: 0.3893; 1.04: 0.62323; 1.05: 0.63; 1.06: 0.6; 1.07: 0.505
| Std Dev. |  | 1.01: 0; 1.02: 0.015055; 1.03: 0.020937; 1.04: 0.0083501; 1.05: 0.013784; 1.06: 0; 1.07: 0.010191
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 2.626%; 1.03: 5.389%; 1.04: 1.34%; 1.05: 2.205%; 1.06: 0%; 1.07: 2.011%
| PDM3 |  | 1.01: -7.974%; 1.02: 5.523%; 1.03: -28.5%; 1.04: 14.69%; 1.05: 15.03%; 1.06: 10.43%; 1.07: -6.747%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0; 1.03: 2.41; 1.04: 0.65; 1.05: 0.67; 1.06: NA; 1.07: 0.87
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for In in OREAS 505b (Certified Value 0.081 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.12 | 1.13 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 3
3
| 1 | 1 | 0.08 | 0.08 | 0.07 | 0.082 | 0.083 | 0.081 | 0.09 | 0.079 | 0.09 | 0.11 | 5.84 | 0.08 | 0.069687 | 0.039 | 0.065 | 0.083 | < 2 | 0.07 | 0.083 | 0.1 | 0.08 | [6 empty] | 1
| 1 | 2 | 0.08 | 0.08 | 0.08 | 0.085 | 0.078 | 0.086 | 0.088 | 0.081 | 0.08 | 0.095 | 8.21 | 0.07 | 0.08561 | 0.038 | 0.069 | 0.084 | < 2 | 0.07 | 0.082 | 0.09 | 0.08 | [6 empty] | 27
| 1 | 3 | 0.09 | 0.08 | 0.08 | 0.083 | 0.094 | 0.087 | 0.093 | 0.086 | 0.09 | 0.087 | 5.58 | 0.08 | 0.08901 | 0.037 | 0.072 | 0.086 | < 2 | 0.07 | 0.082 | 0.1 | 0.08 | [6 empty] | 16
| 1 | 4 | 0.08 | 0.09 | 0.08 | 0.085 | 0.093 | 0.086 | 0.089 | 0.079 | 0.09 | 0.125 | 9.26 | 0.08 | 0.076775 | 0.036 | 0.083 | 0.084 | < 2 | 0.07 | 0.083 | 0.11 | 0.08 | [6 empty] | 0.081453
| 1 | 5 | 0.09 | 0.09 | 0.07 | 0.086 | 0.082 | 0.083 | 0.091 | 0.081 | 0.08 | 0.109 | 8.18 | 0.08 | 0.07136 | 0.041 | 0.082 | 0.095 | < 2 | 0.07 | 0.082 | 0.1 | 0.08 | [6 empty] | 32
| 1 | 6 | 0.07 | 0.08 | 0.07 | 0.082 | 0.094 | 0.082 | 0.087 | 0.08 | 0.09 | 0.123 | 5.46 | 0.08 | 0.092857 | 0.04 | 0.086 | 0.082 | < 2 | 0.07 | 0.087 | 0.09 | 0.08
| Mean |  | 0.081667 | 0.083333 | 0.075 | 0.083833 | 0.087333 | 0.084167 | 0.089667 | 0.081 | 0.086667 | 0.10817 | 7.0883 | 0.078333 | 0.080883 | 0.0385 | 0.076167 | 0.085667 |  | 0.07 | 0.083167 | 0.098333 | 0.08
| Median |  | 0.08 | 0.08 | 0.075 | 0.084 | 0.088 | 0.0845 | 0.0895 | 0.0805 | 0.09 | 0.1095 | 7.01 | 0.08 | 0.081193 | 0.0385 | 0.077 | 0.084 |  | 0.07 | 0.0825 | 0.1 | 0.08
| Std Dev. |  | 0.0075277 | 0.005164 | 0.0054772 | 0.0017224 | 0.0071461 | 0.0024833 | 0.0021602 | 0.0026077 | 0.005164 | 0.015026 | 1.6523 | 0.0040825 | 0.0096403 | 0.0018708 | 0.008612 | 0.004761 |  | 0 | 0.0019408 | 0.0075277 | 0
| Rel.Std.Dev. |  | 9.218% | 6.197% | 7.303% | 2.055% | 8.183% | 2.95% | 2.409% | 3.219% | 5.958% | 13.89% | 23.31% | 5.212% | 11.92% | 4.859% | 11.31% | 5.558% |  | 0% | 2.334% | 7.655% | 0%
| PDM3 |  | 0.2622% | 2.308% | -7.923% | 2.922% | 7.219% | 3.331% | 10.08% | -0.5563% | 6.401% | 32.8% | 8602% | -3.83% | -0.6995% | -52.73% | -6.49% | 5.173% |  | -14.06% | 2.104% | 20.72% | -1.784%
| Z-Score (Absolute) |  | 0.28 | 0 | 1.4 | 0.08 | 0.67 | 0.14 | 1.07 | 0.39 | 0.56 | 4.19 | 1180.88 | 0.84 | 0.41 | 7.56 | 1.21 | 0.39 | 154.53 | 2.25 | 0.03 | 2.53 | 0.56
| Analytical results for K in OREAS 505b (Certified Value 3.19 wt.%) | [28 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 3.15 | 3.25 | 3.19 | 3.31 | 3.11 | 3.15 | 2.83 | 3.14 | 3.2064 | 3.1695 | 3.47 | 4.3 | 3.0009 | 3.26 | 3.27 | 3.2786 | 3.11 | 3.13 | 3.25 | 2.86 | 2.15 | 3.0399 | 3.1288 | 3.17 | 3.13 | 3.23 |  | 1
| 1 | 2 | 3.21 | 3.15 | 3.13 | 3.27 | 3.24 | 3.17 | 2.73 | 3.2 | 3.2168 | 3.1946 | 3.52 | 3.81 | 3.0883 | 3.24 | 3.21 | 3.2312 | 3.08 | 3.18 | 3.38 | 2.87 | 2.55 | 3.0784 | 3.057 | 3.18 | 3.17 | 3.23 |  | #N/A
| 1 | 3 | 3.09 | 2.92 | 3.26 | 3.3 | 3.24 | 3.17 | 3.27 | 3.18 | 3.2988 | 3.1666 | 3.42 | 4.51 | 3.0477 | 3.23 | 3.24 | 3.1965 | 3.12 | 3.17 | 3.23 | 2.81 | 2.59 | 3.0236 | 3.1096 | 3.33 | 3.14 | 3.23 |  | 16
| 1 | 4 | 3.2 | 3.14 | 3.22 | 3.22 | 3.22 | 3.2 | 3.41 | 3.22 | 3.2867 | 3.1076 | 3.5 | 4.81 | 3.0121 | 3.2 | 3.29 | 3.1732 | 3.12 | 3.15 | 3.31 | 2.86 | 2.53 | 3.1167 | 3.1021 | 3.25 | 3.16 | 3.25 |  | 3.1874
| 1 | 5 | 3.2 | 3.09 | 3.27 | 3.26 | 3.32 | 3.19 | 3.2 | 3.11 | 3.2994 | 3.1622 | 3.4 | 3.94 | 3.0256 | 3.23 | 3.23 | 3.2403 | 3.07 | 3.18 | 3.45 | 2.88 | 2.19 | 3.3683 | 3.0548 | 3.19 | 3.2 | 3.2 |  | 33
| 1 | 6 | 3.16 | 2.64 | 3.18 | 3.26 | 3.22 | 3.14 | 3.26 | 3.12 | 3.3066 | 3.1654 | 3.52 | 4.46 | 3.0355 | 3.2 | 3.24 | 3.287 | 3.09 | 3.1 | 3.24 | 2.89 | 1.92 | 3.1739 | 3.1106 | 3.23 | 3.01 | 3.22
| Mean |  | 3.1683 | 3.0317 | 3.2083 | 3.27 | 3.225 | 3.17 | 3.1167 | 3.1617 | 3.2691 | 3.161 | 3.4717 | 4.305 | 3.035 | 3.2267 | 3.2467 | 3.2345 | 3.0983 | 3.1517 | 3.31 | 2.8617 | 2.3217 | 3.1335 | 3.0939 | 3.225 | 3.135 | 3.2267
| Median |  | 3.18 | 3.115 | 3.205 | 3.265 | 3.23 | 3.17 | 3.23 | 3.16 | 3.2927 | 3.166 | 3.485 | 4.38 | 3.0305 | 3.23 | 3.24 | 3.2357 | 3.1 | 3.16 | 3.28 | 2.865 | 2.36 | 3.0975 | 3.1059 | 3.21 | 3.15 | 3.23
| Std Dev. |  | 0.04535 | 0.22031 | 0.052694 | 0.032249 | 0.067454 | 0.022804 | 0.27156 | 0.044907 | 0.045128 | 0.028661 | 0.051543 | 0.37399 | 0.030913 | 0.023381 | 0.028752 | 0.044606 | 0.02137 | 0.031885 | 0.088769 | 0.027869 | 0.27411 | 0.12721 | 0.030662 | 0.059917 | 0.065955 | 0.01633
| Rel.Std.Dev. |  | 1.431% | 7.267% | 1.642% | 0.9862% | 2.092% | 0.7194% | 8.713% | 1.42% | 1.38% | 0.9067% | 1.485% | 8.687% | 1.019% | 0.7246% | 0.8856% | 1.379% | 0.6897% | 1.012% | 2.682% | 0.9739% | 11.81% | 4.06% | 0.9911% | 1.858% | 2.104% | 0.5061%
| PDM3 |  | -0.5989% | -4.887% | 0.6561% | 2.591% | 1.179% | -0.5466% | -2.22% | -0.808% | 2.563% | -0.8288% | 8.918% | 35.06% | -4.781% | 1.231% | 1.859% | 1.476% | -2.795% | -1.122% | 3.846% | -10.22% | -27.16% | -1.693% | -2.935% | 1.179% | -1.645% | 1.231%
| Z-Score (Absolute) |  | 0.01 | 1.51 | 0.43 | 1.11 | 0.61 | 0.01 | 0.58 | 0.08 | 1.1 | 0.09 | 3.32 | 12.47 | 1.47 | 0.63 | 0.85 | 0.72 | 0.78 | 0.19 | 1.55 | 3.38 | 9.31 | 0.39 | 0.83 | 0.61 | 0.38 | 0.63
| Analytical results for La in OREAS 505b (Certified Value 33.6 ppm) | [28 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  |  | 1
2
| 1 | 1 | 34.5 | 31.3 | 33.7 | 34.3 | 33.3 | 35 | 32.2 | 36.15 | 26.663 | 34.7943 | 42.8 | 29.8 | 41.8 | 35 | 36.6 | 33 | 33.4 | 33.3 | 26.704 | 22.436 | 34 | 31.7 | 32.8 |  |  |  |  | 1
| 1 | 2 | 34.2 | 30.1 | 33.83 | 33.6 | 34.1 | 33.8 | 32.8 | 34.57 | 29.7925 | 33.8843 | 42.4 | 29.47 | 39 | 34 | 35.6 | 33.9 | 32.5 | 36.1 | 27.985 | 21.96 | 35 | 27.4 | 33.7 |  |  |  |  | 28
| 1 | 3 | 32.4 | 32 | 33.91 | 33.7 | 33.6 | 35.4 | 35.8 | 35.65 | 28.0176 | 36.0885 | 43.8 | 29.07 | 36.4 | 34 | 32.9 | 32.9 | 32.1 | 33.8 | 26.077 | 22.598 | 38 | 28 | 34.2 |  |  |  |  | 16
| 1 | 4 | 34.6 | 30 | 33.72 | 35.4 | 34.4 | 36.6 | 32.7 | 36.34 | 27.5602 | 34.9657 | 43.2 | 28.71 | 39.7 | 34 | 34 | 34.5 | 32.1 | 34.3 | 26.728 | 22.47 | 36 | 30.6 | 33.7 |  |  |  |  | 33.6182
| 1 | 5 | 32.8 | 30.6 | 33.67 | 35.3 | 34.3 | 33.8 | 31.4 | 37.28 | 29.4343 | 34.6972 | 42.9 | 29.93 | 39.9 | 35 | 35.3 | 34.2 | 31.4 | 32.1 | 26.432 | 21.777 | 37 | 28.4 | 32.9 |  |  |  |  | 34
| 1 | 6 | 33.6 | 30.9 | 35.15 | 35.5 | 32.5 | 34.6 | 33.2 | 34.2 | 28.4599 | 34.8567 | 45 | 28.1 | 43.7 | 34 | 32.6 | 33.1 | 32.1 | 32.8 | 27.1 | 22.96 | 37 | 26.6 | 33.4
| Mean |  | 33.6833 | 30.8167 | 33.9967 | 34.6333 | 33.7 | 34.8667 | 33.0167 | 35.6983 | 28.3212 | 34.8811 | 43.35 | 29.18 | 40.0833 | 34.3333 | 34.5 | 33.6 | 32.2667 | 33.7333 | 26.8377 | 22.3668 | 36.1667 | 28.7833 | 33.45
| Median |  | 33.9 | 30.75 | 33.775 | 34.8 | 33.85 | 34.8 | 32.75 | 35.9 | 28.2387 | 34.8255 | 43.05 | 29.27 | 39.8 | 34 | 34.65 | 33.5 | 32.1 | 33.55 | 26.716 | 22.453 | 36.5 | 28.2 | 33.55
| Std Dev. |  | 0.91742 | 0.75741 | 0.57214 | 0.8756 | 0.72388 | 1.0633 | 1.4972 | 1.1522 | 1.1699 | 0.70719 | 0.93327 | 0.69719 | 2.4879 | 0.5164 | 1.5925 | 0.68702 | 0.65929 | 1.3895 | 0.657 | 0.43223 | 1.472 | 1.9621 | 0.53198
| Rel.Std.Dev. |  | 2.724% | 2.458% | 1.683% | 2.528% | 2.148% | 3.05% | 4.535% | 3.228% | 4.131% | 2.027% | 2.153% | 2.389% | 6.207% | 1.504% | 4.616% | 2.045% | 2.043% | 4.119% | 2.448% | 1.932% | 4.07% | 6.817% | 1.59%
| PDM3 |  | 0.1939% | -8.333% | 1.126% | 3.02% | 0.2434% | 3.714% | -1.789% | 6.188% | -15.76% | 3.757% | 28.95% | -13.2% | 19.23% | 2.127% | 2.623% | -0.05403% | -4.02% | 0.3426% | -20.17% | -33.47% | 7.581% | -14.38% | -0.5002%
| Z-Score (Absolute) |  | 0.01 | 1.65 | 0.17 | 0.53 | 0 | 0.67 | 0.39 | 1.14 | 3.07 | 0.67 | 5.51 | 2.58 | 3.64 | 0.36 | 0.46 | 0.06 | 0.82 | 0.02 | 3.92 | 6.47 | 1.41 | 2.81 | 0.14
| Analytical results for Li in OREAS 505b (Certified Value 47.1 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 48.2 | 49.7 | 41.9 | 47.5 | 45.2 | 47.2 | 48.4 | 48.1 | 45.6 | 48.9594 | 48.4329 | 46.6 | 45.59 | 47.26 | 55 | 51.1615 | 45 | 47.5 | 47.1 | 44.2 | 45.8 | 44.943 | 44.343 | 49.1 | 45 | 45 |  | 1
| 1 | 2 | 48.1 | 49.2 | 42.8 | 51 | 47.2 | 48.7 | 50.1 | 48.2 | 45 | 49.2191 | 49.8118 | 47.4 | 41.41 | 46.98 | 51 | 51.3887 | 46 | 48.1 | 49.1 | 45.2 | 47.4 | 44.113 | 42.452 | 49.7 | 46 | 44 |  | #N/A
| 1 | 3 | 48.7 | 47.6 | 42.3 | 48.9 | 46.5 | 47.4 | 49.2 | 47.5 | 45.7 | 48.4379 | 48.6253 | 45.9 | 43.85 | 47.59 | 49 | 50.2951 | 46 | 47.2 | 48.9 | 43.1 | 47.4 | 45.618 | 43.18 | 51.9 | 46 | 45 |  | 16
| 1 | 4 | 48 | 47.9 | 41.5 | 50.8 | 47.1 | 48.5 | 51.2 | 48.4 | 46.6 | 49.2394 | 47.7138 | 47.1 | 45.48 | 46.32 | 53 | 51.7165 | 46 | 47.6 | 49.8 | 45.8 | 46.6 | 45.181 | 43.533 | 50.6 | 46 | 44 |  | 47.0743
| 1 | 5 | 48.8 | 52.4 | 42.2 | 46.8 | 48.3 | 47.8 | 48.7 | 47 | 45.2 | 49.5515 | 48.1801 | 47.1 | 43.61 | 47.51 | 49 | 50.4776 | 45 | 47.3 | 50.6 | 44.9 | 45 | 47.718 | 42.504 | 48.7 | 45 | 45 |  | 35
| 1 | 6 | 48.9 | 50.1 | 41.9 | 46.1 | 46.9 | 46.6 | 50.2 | 47.1 | 45.7 | 49.1886 | 48.062 | 47.9 | 44.75 | 47.06 | 54 | 51.4318 | 45 | 46 | 47.4 | 43.2 | 45.7 | 45.212 | 44.711 | 46.4 | 44 | 44
| Mean |  | 48.45 | 49.4833 | 42.1 | 48.5167 | 46.8667 | 47.7 | 49.6333 | 47.7167 | 45.6333 | 49.0993 | 48.471 | 47 | 44.115 | 47.12 | 51.8333 | 51.0785 | 45.5 | 47.2833 | 48.8167 | 44.4 | 46.3167 | 45.4642 | 43.4538 | 49.4 | 45.3333 | 44.5
| Median |  | 48.45 | 49.45 | 42.05 | 48.2 | 47 | 47.6 | 49.65 | 47.8 | 45.65 | 49.2039 | 48.3065 | 47.1 | 44.3 | 47.16 | 52 | 51.2751 | 45.5 | 47.4 | 49 | 44.55 | 46.2 | 45.1965 | 43.3565 | 49.4 | 45.5 | 44.5
| Std Dev. |  | 0.3937 | 1.736 | 0.44272 | 2.0663 | 1.0132 | 0.8 | 1.0558 | 0.59805 | 0.55377 | 0.37508 | 0.7278 | 0.68702 | 1.5546 | 0.45952 | 2.5626 | 0.56743 | 0.54772 | 0.70261 | 1.3556 | 1.0973 | 0.98065 | 1.212 | 0.93362 | 1.8633 | 0.8165 | 0.54772
| Rel.Std.Dev. |  | 0.8126% | 3.508% | 1.052% | 4.259% | 2.162% | 1.677% | 2.127% | 1.253% | 1.214% | 0.7639% | 1.502% | 1.462% | 3.524% | 0.9752% | 4.944% | 1.111% | 1.204% | 1.486% | 2.777% | 2.471% | 2.117% | 2.666% | 2.149% | 3.772% | 1.801% | 1.231%
| PDM3 |  | 2.922% | 5.117% | -10.57% | 3.064% | -0.4412% | 1.329% | 5.436% | 1.364% | -3.061% | 4.302% | 2.967% | -0.1579% | -6.287% | 0.09699% | 10.11% | 8.506% | -3.344% | 0.444% | 3.701% | -5.681% | -1.61% | -3.42% | -7.691% | 4.94% | -3.698% | -5.469%
| Z-Score (Absolute) |  | 0.49 | 0.89 | 2 | 0.52 | 0.13 | 0.2 | 0.95 | 0.2 | 0.61 | 0.74 | 0.5 | 0.08 | 1.21 | 0.03 | 1.82 | 1.52 | 0.67 | 0.03 | 0.63 | 1.1 | 0.35 | 0.68 | 1.47 | 0.86 | 0.73 | 1.06
| Analytical results for Lu in OREAS 505b (Certified Value 0.19 ppm) | [28 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.17; 1.06: 1.19; 1.07: 1.23; 1.08: 1.26; 1.09: 1.29 |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS |  | 2
3
| 1 | 1 | 1.01: 0.1; 1.02: 0.18; 1.03: 0.1191; 1.04: 0.20248; 1.05: 0.2; 1.06: 0.19; 1.07: 0.2; 1.08: 0.255; 1.09: 0.19 |  | 1
| 1 | 2 | 1.01: 0.1; 1.02: 0.18; 1.03: 0.1295; 1.04: 0.19361; 1.05: 0.19; 1.06: 0.19; 1.07: 0.2; 1.08: 0.238; 1.09: 0.18 |  | 29
| 1 | 3 | 1.01: 0.2; 1.02: 0.19; 1.03: 0.1169; 1.04: 0.19477; 1.05: 0.19; 1.06: 0.17; 1.07: 0.2; 1.08: 0.245; 1.09: 0.17 |  | 16
| 1 | 4 | 1.01: 0.1; 1.02: 0.17; 1.03: 0.1176; 1.04: 0.19902; 1.05: 0.19; 1.06: 0.19; 1.07: 0.2; 1.08: 0.24; 1.09: 0.18 |  | 0.18721
| 1 | 5 | 1.01: 0.1; 1.02: 0.18; 1.03: 0.1223; 1.04: 0.1994; 1.05: 0.19; 1.06: 0.19; 1.07: 0.2; 1.08: 0.245; 1.09: 0.17 |  | 36
| 1 | 6 | 1.01: 0.1; 1.02: 0.19; 1.03: 0.1258; 1.04: 0.19706; 1.05: 0.21; 1.06: 0.17; 1.07: 0.2; 1.08: 0.244; 1.09: 0.18
| Mean |  | 1.01: 0.11667; 1.02: 0.18167; 1.03: 0.12187; 1.04: 0.19772; 1.05: 0.195; 1.06: 0.18333; 1.07: 0.2; 1.08: 0.2445; 1.09: 0.17833
| Median |  | 1.01: 0.1; 1.02: 0.18; 1.03: 0.1207; 1.04: 0.19804; 1.05: 0.19; 1.06: 0.19; 1.07: 0.2; 1.08: 0.2445; 1.09: 0.18
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.0075277; 1.03: 0.0049891; 1.04: 0.0032628; 1.05: 0.0083666; 1.06: 0.010328; 1.07: 3.0405e-17; 1.08: 0.0058907; 1.09: 0.0075277
| Rel.Std.Dev. |  | 1.01: 34.99%; 1.02: 4.144%; 1.03: 4.094%; 1.04: 1.65%; 1.05: 4.291%; 1.06: 5.633%; 1.07: 1.52e-14%; 1.08: 2.409%; 1.09: 4.221%
| PDM3 |  | 1.01: -37.68%; 1.02: -2.962%; 1.03: -34.9%; 1.04: 5.615%; 1.05: 4.16%; 1.06: -2.072%; 1.07: 6.831%; 1.08: 30.6%; 1.09: -4.742%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.1; 1.03: 3.55; 1.04: 0.83; 1.05: 0.67; 1.06: 0; 1.07: NA; 1.08: 3.54; 1.09: 0.29
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 505b (Certified Value 0.789 wt.%) | [28 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.812 | 0.84 | 0.76 | 0.82 | 0.75 | 0.75 | 0.78 | 0.79 | 0.7946 | 0.77409 | 0.8 | 1.01 | 0.89636 | 0.81 | 0.78 | 0.83182 | 0.79 | 0.75 | 0.76 | 0.81 | 0.8 | 0.85811 | 0.82651 | 0.73 | 0.74 | 0.734 |  | 1
| 1 | 2 | 0.814 | 0.81 | 0.77 | 0.81 | 0.77 | 0.77 | 0.79 | 0.8 | 0.8047 | 0.77002 | 0.83 | 0.88 | 0.88424 | 0.81 | 0.77 | 0.81316 | 0.78 | 0.77 | 0.79 | 0.8 | 0.81 | 0.84065 | 0.80044 | 0.72 | 0.77 | 0.734 |  | #N/A
| 1 | 3 | 0.809 | 0.81 | 0.77 | 0.81 | 0.77 | 0.76 | 0.78 | 0.77 | 0.8223 | 0.79076 | 0.81 | 0.89 | 0.8913 | 0.81 | 0.76 | 0.80885 | 0.79 | 0.76 | 0.79 | 0.82 | 0.8 | 0.83879 | 0.81989 | 0.77 | 0.74 | 0.741 |  | 16
| 1 | 4 | 0.81 | 0.84 | 0.76 | 0.8 | 0.77 | 0.76 | 0.82 | 0.8 | 0.8225 | 0.76896 | 0.8 | 1.11 | 0.89335 | 0.8 | 0.77 | 0.80046 | 0.8 | 0.76 | 0.8 | 0.81 | 0.79 | 0.85079 | 0.82223 | 0.76 | 0.75 | 0.741 |  | 0.78878
| 1 | 5 | 0.832 | 0.8 | 0.78 | 0.8 | 0.79 | 0.75 | 0.77 | 0.78 | 0.8249 | 0.77748 | 0.8 | 0.92 | 0.89096 | 0.81 | 0.77 | 0.81985 | 0.79 | 0.76 | 0.82 | 0.8 | 0.8 | 0.85652 | 0.80418 | 0.71 | 0.75 | 0.733 |  | 37
| 1 | 6 | 0.824 | 0.84 | 0.77 | 0.81 | 0.77 | 0.75 | 0.79 | 0.76 | 0.8242 | 0.78495 | 0.79 | 0.94 | 0.89244 | 0.8 | 0.77 | 0.82775 | 0.79 | 0.74 | 0.77 | 0.8 | 0.8 | 0.85432 | 0.81809 | 0.74 | 0.71 | 0.734
| Mean |  | 0.81683 | 0.82333 | 0.76833 | 0.80833 | 0.77 | 0.75667 | 0.78833 | 0.78333 | 0.81553 | 0.77771 | 0.805 | 0.95833 | 0.89144 | 0.80667 | 0.77 | 0.81698 | 0.79 | 0.75667 | 0.78833 | 0.80667 | 0.8 | 0.84986 | 0.81522 | 0.73833 | 0.74333 | 0.73617
| Median |  | 0.813 | 0.825 | 0.77 | 0.81 | 0.77 | 0.755 | 0.785 | 0.785 | 0.8224 | 0.77579 | 0.8 | 0.93 | 0.89187 | 0.81 | 0.77 | 0.8165 | 0.79 | 0.76 | 0.79 | 0.805 | 0.8 | 0.85255 | 0.81899 | 0.735 | 0.745 | 0.734
| Std Dev. |  | 0.0091742 | 0.018619 | 0.0075277 | 0.0075277 | 0.012649 | 0.008165 | 0.017224 | 0.01633 | 0.012749 | 0.0086208 | 0.013784 | 0.087502 | 0.0040252 | 0.005164 | 0.0063246 | 0.011819 | 0.0063246 | 0.010328 | 0.02137 | 0.008165 | 0.0063246 | 0.0082533 | 0.010459 | 0.023166 | 0.019664 | 0.0037639
| Rel.Std.Dev. |  | 1.123% | 2.261% | 0.9797% | 0.9313% | 1.643% | 1.079% | 2.185% | 2.085% | 1.563% | 1.108% | 1.712% | 9.131% | 0.4515% | 0.6402% | 0.8214% | 1.447% | 0.8006% | 1.365% | 2.711% | 1.012% | 0.7906% | 0.9711% | 1.283% | 3.138% | 2.645% | 0.5113%
| PDM3 |  | 3.556% | 4.38% | -2.593% | 2.478% | -2.381% | -4.072% | -0.05719% | -0.6911% | 3.391% | -1.404% | 2.056% | 21.49% | 13.01% | 2.267% | -2.381% | 3.575% | 0.1541% | -4.072% | -0.05719% | 2.267% | 1.422% | 7.743% | 3.351% | -6.396% | -5.762% | -6.671%
| Z-Score (Absolute) |  | 0.67 | 0.87 | 0.82 | 0.41 | 0.77 | 1.18 | 0.21 | 0.36 | 0.63 | 0.53 | 0.31 | 5.03 | 2.97 | 0.36 | 0.77 | 0.68 | 0.15 | 1.18 | 0.21 | 0.36 | 0.15 | 1.69 | 0.62 | 1.74 | 1.59 | 1.81
| Analytical results for Mn in OREAS 505b (Certified Value 0.042 wt.%) | [28 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.0399 | 0.0445 | 0.0412 | 0.0433 | 0.0406 | 0.0414 | 0.0419 | 0.0425 | 0.0415 | 0.043773 | 0.0416 | 0.039 | 0.043957 | 0.04296 | 0.0432 | 0.042472 | 0.0417 | 0.0418 | 0.0414 | 0.0409 | 0.0433 | 0.041755 | 0.038837 | 0.0428 | 0.0442 | 0.0394 |  | 1
| 1 | 2 | 0.0411 | 0.0428 | 0.0419 | 0.0427 | 0.0418 | 0.0421 | 0.0424 | 0.0428 | 0.0421 | 0.043564 | 0.0432 | 0.04 | 0.044337 | 0.04317 | 0.0427 | 0.043094 | 0.041 | 0.0425 | 0.0428 | 0.0403 | 0.0436 | 0.042042 | 0.037654 | 0.0445 | 0.0412 | 0.0396 |  | 30
| 1 | 3 | 0.0397 | 0.044 | 0.0418 | 0.0427 | 0.0408 | 0.0416 | 0.0419 | 0.0415 | 0.0431 | 0.04187 | 0.0421 | 0.036 | 0.044174 | 0.04298 | 0.043 | 0.043195 | 0.0418 | 0.042 | 0.0425 | 0.0402 | 0.0426 | 0.041562 | 0.038297 | 0.0454 | 0.0408 | 0.0397 |  | 16
| 1 | 4 | 0.0408 | 0.0448 | 0.0411 | 0.0422 | 0.0414 | 0.0419 | 0.0436 | 0.0435 | 0.0429 | 0.041102 | 0.042 | 0.04 | 0.044735 | 0.04251 | 0.0432 | 0.043431 | 0.0418 | 0.0423 | 0.0435 | 0.0408 | 0.0419 | 0.042151 | 0.038513 | 0.0423 | 0.0432 | 0.0397 |  | 0.042298
| 1 | 5 | 0.0411 | 0.0434 | 0.0423 | 0.0424 | 0.0429 | 0.0417 | 0.0409 | 0.042 | 0.0432 | 0.041632 | 0.0418 | 0.039 | 0.044344 | 0.04258 | 0.0432 | 0.042892 | 0.042 | 0.0423 | 0.0438 | 0.0401 | 0.0418 | 0.042508 | 0.037759 | 0.0449 | 0.0423 | 0.039 |  | 38
| 1 | 6 | 0.0407 | 0.0455 | 0.0413 | 0.0425 | 0.0414 | 0.041 | 0.0426 | 0.0414 | 0.0433 | 0.041368 | 0.0419 | 0.038 | 0.044251 | 0.04248 | 0.0439 | 0.042294 | 0.0423 | 0.0418 | 0.0421 | 0.0406 | 0.0419 | 0.042694 | 0.03819 | 0.0485 | 0.0412 | 0.0394
| Mean |  | 0.04055 | 0.044167 | 0.0416 | 0.042633 | 0.041483 | 0.041617 | 0.042217 | 0.042283 | 0.042683 | 0.042218 | 0.0421 | 0.038667 | 0.0443 | 0.04278 | 0.0432 | 0.042896 | 0.041767 | 0.042117 | 0.042683 | 0.040483 | 0.042517 | 0.042119 | 0.038208 | 0.044733 | 0.04215 | 0.039467
| Median |  | 0.04075 | 0.04425 | 0.04155 | 0.0426 | 0.0414 | 0.04165 | 0.04215 | 0.04225 | 0.043 | 0.041751 | 0.04195 | 0.039 | 0.044294 | 0.04277 | 0.0432 | 0.042993 | 0.0418 | 0.04215 | 0.04265 | 0.04045 | 0.04225 | 0.042096 | 0.038244 | 0.0447 | 0.04175 | 0.0395
| Std Dev. |  | 0.00060581 | 0.00097707 | 0.00047329 | 0.00037771 | 0.00082077 | 0.00038687 | 0.00089759 | 0.0008085 | 0.00072226 | 0.0011542 | 0.00056569 | 0.0015055 | 0.0002563 | 0.00029237 | 0.00039497 | 0.00043719 | 0.00043205 | 0.00029269 | 0.00088863 | 0.00033116 | 0.00078337 | 0.00043158 | 0.00044856 | 0.0022079 | 0.0013383 | 0.00026583
| Rel.Std.Dev. |  | 1.494% | 2.212% | 1.138% | 0.886% | 1.979% | 0.9296% | 2.126% | 1.912% | 1.692% | 2.734% | 1.344% | 3.894% | 0.5786% | 0.6834% | 0.9143% | 1.019% | 1.034% | 0.6949% | 2.082% | 0.818% | 1.842% | 1.025% | 1.174% | 4.936% | 3.175% | 0.6736%
| PDM3 |  | -4.133% | 4.417% | -1.651% | 0.7921% | -1.927% | -1.612% | -0.193% | -0.03539% | 0.9103% | -0.1894% | -0.4688% | -8.586% | 4.732% | 1.139% | 2.132% | 1.414% | -1.257% | -0.4294% | 0.9103% | -4.291% | 0.5162% | -0.4247% | -9.669% | 5.757% | -0.3506% | -6.694%
| Z-Score (Absolute) |  | 1.92 | 2.33 | 0.68 | 0.53 | 0.82 | 0.66 | 0.04 | 0.12 | 0.59 | 0.04 | 0.1 | 4.12 | 2.48 | 0.7 | 1.19 | 0.84 | 0.49 | 0.08 | 0.59 | 1.99 | 0.39 | 0.08 | 4.66 | 2.99 | 0.04 | 3.19
| Analytical results for Mo in OREAS 505b (Certified Value 68 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 70.2 | 67.2 | 63.45 | 73.29 | 68.5 | 69 | 71.3 | 70 | 69.1 | 73.53 | 68.0288 | 76.3 | 59.72 | 68.3 | 67.8 | 67.5 | 67.7272 | 71 | 68.8 | 68.9 | 69.6 | 68.21 | 66.74 | 64.734 | 77.9 | 70.15 | 63 | 1
| 1 | 2 | 69.4 | 65.2 | 64.86 | 69.02 | 70.7 | 70 | 73.1 | 71.2 | 68.8 | 74.18 | 68.2207 | 75.4 | 61.77 | 68.95 | 68.23 | 61.6 | 66.8891 | 68 | 61.9 | 68.8 | 69 | 68.02 | 66.58 | 64.256 | 78 | 68.32 | 64 | 31
| 1 | 3 | 70 | 67.3 | 64.02 | 69.85 | 69.6 | 68.9 | 75 | 71 | 69.7 | 74.65 | 67.3728 | 77.6 | 62.28 | 71.55 | 68.55 | 61.6 | 66.9764 | 68 | 65.2 | 67 | 68.5 | 66.79 | 66.85 | 64.54 | 77.5 | 66.33 | 63 | 16
| 1 | 4 | 70.4 | 68.5 | 63.2 | 71.19 | 72.8 | 68.6 | 76.4 | 70.8 | 69.6 | 75.52 | 66.9402 | 76.4 | 64.24 | 71.46 | 68.63 | 67 | 68.5734 | 68 | 71.5 | 70.2 | 68.6 | 66.69 | 67.88 | 65.142 | 78.1 | 70.11 | 65 | 68.1368
| 1 | 5 | 69.5 | 66 | 63.14 | 68.44 | 73 | 68.1 | 73.1 | 69.8 | 69.4 | 74.95 | 67.0227 | 74.7 | 64.01 | 72.11 | 68.22 | 62.8 | 66.1649 | 68 | 67.1 | 70.8 | 67.6 | 66.88 | 68.4 | 64.728 | 78.6 | 71.01 | 64 | 39
| 1 | 6 | 69.2 | 69.1 | 64.33 | 67.55 | 70.9 | 64.8 | 73.8 | 70.3 | 70.4 | 73.8 | 67.8482 | 75.6 | 62.26 | 71.45 | 67.82 | 67.8 | 68.2267 | 66 | 70 | 68.5 | 68.9 | 66.98 | 67.86 | 65.224 | 78.1 | 67.57 | 63
| Mean |  | 69.7833 | 67.2167 | 63.8333 | 69.89 | 70.9167 | 68.2333 | 73.7833 | 70.5167 | 69.5 | 74.4383 | 67.5722 | 76 | 62.38 | 70.6367 | 68.2083 | 64.7167 | 67.4263 | 68.1667 | 67.4167 | 69.0333 | 68.7 | 67.2617 | 67.385 | 64.7707 | 78.0333 | 68.915 | 63.6667
| Median |  | 69.75 | 67.25 | 63.735 | 69.435 | 70.8 | 68.75 | 73.45 | 70.55 | 69.5 | 74.415 | 67.6105 | 75.95 | 62.27 | 71.455 | 68.225 | 64.9 | 67.3518 | 68 | 67.95 | 68.85 | 68.75 | 66.93 | 67.355 | 64.731 | 78.05 | 69.215 | 63.5
| Std Dev. |  | 0.48339 | 1.4662 | 0.68765 | 2.0778 | 1.7612 | 1.7941 | 1.7566 | 0.56716 | 0.55136 | 0.7451 | 0.53784 | 1.002 | 1.6486 | 1.5905 | 0.35005 | 3.0189 | 0.90875 | 1.6021 | 3.4845 | 1.3397 | 0.66332 | 0.67062 | 0.75514 | 0.3646 | 0.3559 | 1.7993 | 0.8165
| Rel.Std.Dev. |  | 0.6927% | 2.181% | 1.077% | 2.973% | 2.483% | 2.629% | 2.381% | 0.8043% | 0.7933% | 1.001% | 0.796% | 1.318% | 2.643% | 2.252% | 0.5132% | 4.665% | 1.348% | 2.35% | 5.169% | 1.941% | 0.9655% | 0.997% | 1.121% | 0.5629% | 0.4561% | 2.611% | 1.282%
| PDM3 |  | 2.417% | -1.35% | -6.316% | 2.573% | 4.08% | 0.1417% | 8.287% | 3.493% | 2.001% | 9.248% | -0.8285% | 11.54% | -8.449% | 3.669% | 0.105% | -5.019% | -1.043% | 0.04387% | -1.057% | 1.316% | 0.8266% | -1.284% | -1.103% | -4.94% | 14.52% | 1.142% | -6.56%
| Z-Score (Absolute) |  | 0.67 | 0.44 | 1.91 | 0.72 | 1.17 | 0 | 2.41 | 0.99 | 0.55 | 2.7 | 0.29 | 3.38 | 2.55 | 1.05 | 0.01 | 1.53 | 0.35 | 0.03 | 0.36 | 0.35 | 0.2 | 0.42 | 0.37 | 1.51 | 4.26 | 0.3 | 1.99
| Analytical results for Na in OREAS 505b (Certified Value 2.11 wt.%) | [28 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 2.11 | 2.078 | 2.051 | 2.15 | 2.04 | 2.06 | 2.09 | 2.07 | 2.0861 | 2.0856 | 2.28 | 3.51 | 2.1473 | 2.06 | 2.08 | 2.182 | 2.09 | 2.09 | 2.2 | 1.93 | 2.091 | 2.0235 | 2.3443 | 2.12 | 2.19 | 2.07 |  | 1
| 1 | 2 | 2.1 | 2.025 | 2.073 | 2.12 | 2.13 | 2.12 | 2.15 | 2.09 | 2.1097 | 2.1222 | 2.38 | 2.02 | 2.118 | 2.09 | 2.05 | 2.1029 | 2.09 | 2.12 | 2.31 | 1.92 | 2.111 | 2.0406 | 2.2918 | 2.17 | 2.21 | 2.066 |  | #N/A
| 1 | 3 | 2.11 | 2.038 | 2.095 | 2.14 | 2.12 | 2.09 | 2.12 | 2.09 | 2.167 | 2.0818 | 2.28 | 3.84 | 2.1258 | 2.06 | 2.09 | 2.1028 | 2.12 | 2.11 | 2.27 | 1.9 | 2.091 | 1.9931 | 2.3392 | 2.22 | 2.19 | 2.056 |  | 16
| 1 | 4 | 2.12 | 2.07 | 2.071 | 2.11 | 2.11 | 2.12 | 2.21 | 2.1 | 2.1571 | 2.0651 | 2.32 | 2.92 | 2.1004 | 2.1 | 2.09 | 2.127 | 2.1 | 2.12 | 2.32 | 1.93 | 2.071 | 2.0718 | 2.3255 | 2.16 | 2.21 | 2.07 |  | 2.1056
| 1 | 5 | 2.14 | 1.998 | 2.082 | 2.12 | 2.16 | 2.1 | 2.1 | 2.05 | 2.1656 | 2.0554 | 2.29 | 3.87 | 2.1107 | 2.08 | 2.04 | 2.1921 | 2.09 | 2.11 | 2.33 | 1.98 | 2.101 | 5.2157 | 2.286 | 2.12 | 2.23 | 2.035 |  | 40
| 1 | 6 | 2.16 | 2.097 | 2.088 | 2.12 | 2.11 | 2.04 | 2.16 | 2.08 | 2.1634 | 2.08 | 2.25 | 1.35 | 2.101 | 2.09 | 2.09 | 2.2279 | 2.09 | 2.07 | 2.21 | 1.96 | 2.091 | 2.0941 | 2.3313 | 2.17 | 2.08 | 2.052
| Mean |  | 2.1233 | 2.051 | 2.0767 | 2.1267 | 2.1117 | 2.0883 | 2.1383 | 2.08 | 2.1415 | 2.0817 | 2.3 | 2.9183 | 2.1172 | 2.08 | 2.0733 | 2.1558 | 2.0967 | 2.1033 | 2.2733 | 1.9367 | 2.0927 | 2.5731 | 2.3197 | 2.16 | 2.185 | 2.0582
| Median |  | 2.115 | 2.054 | 2.0775 | 2.12 | 2.115 | 2.095 | 2.135 | 2.085 | 2.1603 | 2.0809 | 2.285 | 3.215 | 2.1143 | 2.085 | 2.085 | 2.1545 | 2.09 | 2.11 | 2.29 | 1.93 | 2.091 | 2.0562 | 2.3284 | 2.165 | 2.2 | 2.061
| Std Dev. |  | 0.022509 | 0.037041 | 0.015475 | 0.015055 | 0.039707 | 0.032506 | 0.04446 | 0.017889 | 0.03474 | 0.022908 | 0.045166 | 1.0365 | 0.017706 | 0.016733 | 0.022509 | 0.052234 | 0.012111 | 0.019664 | 0.056804 | 0.028752 | 0.013292 | 1.2951 | 0.02477 | 0.037417 | 0.053572 | 0.013571
| Rel.Std.Dev. |  | 1.06% | 1.806% | 0.7452% | 0.7079% | 1.88% | 1.557% | 2.079% | 0.86% | 1.622% | 1.1% | 1.964% | 35.52% | 0.8363% | 0.8045% | 1.086% | 2.423% | 0.5776% | 0.9349% | 2.499% | 1.485% | 0.6352% | 50.33% | 1.068% | 1.732% | 2.452% | 0.6594%
| PDM3 |  | 0.8413% | -2.594% | -1.375% | 0.9996% | 0.2872% | -0.8209% | 1.554% | -1.217% | 1.703% | -1.136% | 9.232% | 38.6% | 0.55% | -1.217% | -1.533% | 2.382% | -0.4252% | -0.1086% | 7.965% | -8.024% | -0.6151% | 22.2% | 10.17% | 2.583% | 3.77% | -2.254%
| Z-Score (Absolute) |  | 0.17 | 1.18 | 0.71 | 0.23 | 0.05 | 0.49 | 0.45 | 0.64 | 0.51 | 0.61 | 3.47 | 15.02 | 0.05 | 0.64 | 0.77 | 0.77 | 0.33 | 0.21 | 2.97 | 3.32 | 0.41 | 8.57 | 3.83 | 0.85 | 1.32 | 1.05
| Analytical results for Nb in OREAS 505b (Certified Value 11.9 ppm) | [28 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.29 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 11.6 | 13.1 | 10.42 | 10.9 | 11.7 | 11.4 | 13.3 | 12.1 | 12.47 | 13.3726 | 32.9976 | 14.1 | 11.8 | 12.5391 | 11.1 | 12 | 12.1 | 10.6 | 12.16 | 16.331 | 12.3 | 12.2 | [5 empty] | 1
| 1 | 2 | 11.7 | 12.6 | 10.64 | 12.4 | 12.7 | 12.1 | 14.1 | 11.9 | 11.98 | 12.4157 | 33.2173 | 14.2 | 11 | 12.3058 | 11.1 | 11.5 | 12.1 | 10.6 | 11.69 | 15.844 | 12.4 | 11.6 | [5 empty] | 6
| 1 | 3 | 11.5 | 12.7 | 10.49 | 12.1 | 12.1 | 12.2 | 14.1 | 12 | 12.26 | 11.5147 | 33.0528 | 15.9 | 11.8 | 12.1959 | 11.3 | 11.3 | 11.6 | 10.6 | 11.21 | 16.03 | 12.9 | 11.8 | [5 empty] | 16
| 1 | 4 | 11.8 | 12.8 | 10.34 | 12.4 | 12.6 | 12.4 | 14.2 | 12 | 12.58 | 15.1616 | 33.7467 | 14.4 | 11.8 | 12.5589 | 10.8 | 11.9 | 12.1 | 10.9 | 11.69 | 16.144 | 12.4 | 12.1 | [5 empty] | 11.8887
| 1 | 5 | 11.8 | 13.7 | 10.41 | 11.8 | 12.8 | 12.2 | 13.5 | 11.7 | 13 | 13.2956 | 32.2819 | 14.8 | 11.1 | 12.4156 | 10.8 | 12.1 | 12.4 | 10.5 | 11.4 | 15.849 | 12.5 | 12.5 | [5 empty] | 41
| 1 | 6 | 12.3 | 13.1 | 10.54 | 10.7 | 12.2 | 11.1 | 13.9 | 11.9 | 12.69 | 12.6796 | 32.9249 | 13.6 | 11.9 | 12.6922 | 11.1 | 12.1 | 12 | 10.9 | 11.31 | 16.387 | 12.5 | 11.8
| Mean |  | 11.7833 | 13 | 10.4733 | 11.7167 | 12.35 | 11.9 | 13.85 | 11.9333 | 12.4967 | 13.0733 | 33.0369 | 14.5 | 11.5667 | 12.4513 | 11.0333 | 11.8167 | 12.05 | 10.6833 | 11.5767 | 16.0975 | 12.5 | 12
| Median |  | 11.75 | 12.95 | 10.455 | 11.95 | 12.4 | 12.15 | 14 | 11.95 | 12.525 | 12.9876 | 33.0252 | 14.3 | 11.8 | 12.4774 | 11.1 | 11.95 | 12.1 | 10.6 | 11.545 | 16.087 | 12.45 | 11.95
| Std Dev. |  | 0.27869 | 0.4 | 0.1069 | 0.74677 | 0.42308 | 0.52154 | 0.36742 | 0.13663 | 0.35251 | 1.2261 | 0.47347 | 0.78994 | 0.40332 | 0.18166 | 0.19664 | 0.33714 | 0.25884 | 0.17224 | 0.3472 | 0.23284 | 0.20976 | 0.32863
| Rel.Std.Dev. |  | 2.365% | 3.077% | 1.021% | 6.374% | 3.426% | 4.383% | 2.653% | 1.145% | 2.821% | 9.379% | 1.433% | 5.448% | 3.487% | 1.459% | 1.782% | 2.853% | 2.148% | 1.612% | 2.999% | 1.446% | 1.678% | 2.739%
| PDM3 |  | -0.8864% | 9.347% | -11.91% | -1.447% | 3.88% | 0.09491% | 16.5% | 0.3753% | 5.114% | 9.964% | 177.9% | 21.96% | -2.709% | 4.732% | -7.195% | -0.606% | 1.357% | -10.14% | -2.625% | 35.4% | 5.142% | 0.936%
| Z-Score (Absolute) |  | 0.35 | 1.41 | 2.25 | 0.45 | 0.47 | 0.18 | 2.65 | 0.13 | 0.68 | 1.52 | 30.47 | 3.59 | 0.66 | 0.62 | 1.44 | 0.3 | 0.04 | 1.95 | 0.65 | 5.91 | 0.69 | 0.04
| Analytical results for Nd in OREAS 505b (Certified Value 28.4 ppm) | [28 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | [5 empty] | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.19; 1.05: 1.23; 1.06: 1.26 | [5 empty] | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*OES/MS | [5 empty] | 1
1
| 1 | 1 | 1.01: 28.9; 1.02: 32.2; 1.03: 24.7062; 1.04: 32.6; 1.05: 29.1; 1.06: 21.391 | [5 empty] | 1
| 1 | 2 | 1.01: 28.5; 1.02: 32; 1.03: 28.4412; 1.04: 32.6; 1.05: 31.5; 1.06: 20.943 | [5 empty] | 7
| 1 | 3 | 1.01: 29; 1.02: 33.4; 1.03: 26.0102; 1.04: 30.7; 1.05: 29.5; 1.06: 21.326 | [5 empty] | 16
| 1 | 4 | 1.01: 28.3; 1.02: 32.6; 1.03: 25.8155; 1.04: 31.5; 1.05: 30.1; 1.06: 21.109 | [5 empty] | 28.3643
| 1 | 5 | 1.01: 28.8; 1.02: 31.6; 1.03: 27.2386; 1.04: 31.3; 1.05: 28.3; 1.06: 20.554 | [5 empty] | 42
| 1 | 6 | 1.01: 28.9; 1.02: 34; 1.03: 26.9343; 1.04: 30.6; 1.05: 29; 1.06: 21.645
| Mean |  | 1.01: 28.7333; 1.02: 32.6333; 1.03: 26.5243; 1.04: 31.55; 1.05: 29.5833; 1.06: 21.1613
| Median |  | 1.01: 28.85; 1.02: 32.4; 1.03: 26.4722; 1.04: 31.4; 1.05: 29.3; 1.06: 21.2175
| Std Dev. |  | 1.01: 0.27325; 1.02: 0.90701; 1.03: 1.2981; 1.04: 0.88261; 1.05: 1.1107; 1.06: 0.3829
| Rel.Std.Dev. |  | 1.01: 0.951%; 1.02: 2.779%; 1.03: 4.894%; 1.04: 2.797%; 1.05: 3.754%; 1.06: 1.809%
| PDM3 |  | 1.01: 1.301%; 1.02: 15.05%; 1.03: -6.487%; 1.04: 11.23%; 1.05: 4.298%; 1.06: -25.39%
| Z-Score (Absolute) |  | 1.01: 0.11; 1.02: 0.93; 1.03: 0.71; 1.04: 0.64; 1.05: 0.11; 1.06: 2.15
| Analytical results for Ni in OREAS 505b (Certified Value 40.1 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 1
2
| 1 | 1 | 40 | 40.2 | 40.3 | 42 | 40.7 | 39.9 | 40.4 | 41.5 | 40.5 | 37.8214 | 39.5549 | 32.7 | 26.65 | 41.32 | 39.18 | 51 | 43.1478 | 39.4 | 39.9 | 40 | 39.5 | 39.1 | 41.994 | 21.83 | 39.3 | 39.8 | 34 | 1
| 1 | 2 | 40 | 39.1 | 40.5 | 42 | 41 | 41.4 | 42 | 40.8 | 40.2 | 38.5719 | 39.5988 | 34.4 | 24.95 | 41.23 | 40.94 | 38.1 | 43.1464 | 38.2 | 36.8 | 40.8 | 39.5 | 37.6 | 41.218 | 20.641 | 38.5 | 38.3 | 35 | 17
| 1 | 3 | 40 | 39.7 | 40.2 | 41 | 41 | 40.2 | 43.5 | 40.8 | 40.5 | 38.0431 | 37.7951 | 33.4 | 29.167 | 41.84 | 38.79 | 40.6 | 42.5307 | 37.4 | 37 | 39.3 | 39.3 | 36.6 | 41.641 | 20.743 | 40.4 | 40.2 | 35 | 16
| 1 | 4 | 39 | 40.1 | 40 | 41 | 42.6 | 40.5 | 42.7 | 41.1 | 41.8 | 37.9372 | 37.8903 | 34 | 27.59 | 41.37 | 39.07 | 44.2 | 43.4611 | 37.3 | 40.4 | 40.8 | 39.3 | 37.3 | 42.187 | 22.173 | 38.7 | 38.9 | 34 | 40.1052
| 1 | 5 | 40 | 39.1 | 40.1 | 41 | 42.4 | 41 | 41.4 | 39.7 | 40.7 | 38.4906 | 38.2187 | 33.1 | 27.74 | 40.55 | 40.13 | 38.8 | 42.7477 | 37.8 | 38 | 41.4 | 38.8 | 38.2 | 42.396 | 21.013 | 39.3 | 38.9 | 34 | 43
| 1 | 6 | 39 | 40.8 | 40.6 | 41 | 41.6 | 38.4 | 44.4 | 41.5 | 41 | 37.1571 | 38.3131 | 33.5 | 27.47 | 42.06 | 38.85 | 40.6 | 43.7719 | 36.6 | 40.1 | 40 | 39.9 | 37.9 | 42.427 | 21.893 | 39.3 | 40.4 | 35
| Mean |  | 39.6667 | 39.8333 | 40.2833 | 41.3333 | 41.55 | 40.2333 | 42.4 | 40.9 | 40.7833 | 38.0035 | 38.5618 | 33.5167 | 27.2612 | 41.395 | 39.4933 | 42.2167 | 43.1343 | 37.7833 | 38.7 | 40.3833 | 39.3833 | 37.7833 | 41.9772 | 21.3822 | 39.25 | 39.4167 | 34.5
| Median |  | 40 | 39.9 | 40.25 | 41 | 41.3 | 40.35 | 42.35 | 40.95 | 40.6 | 37.9901 | 38.2659 | 33.45 | 27.53 | 41.345 | 39.125 | 40.6 | 43.1471 | 37.6 | 38.95 | 40.4 | 39.4 | 37.75 | 42.0905 | 21.4215 | 39.3 | 39.35 | 34.5
| Std Dev. |  | 0.5164 | 0.66833 | 0.23166 | 0.5164 | 0.79436 | 1.0482 | 1.4464 | 0.66633 | 0.56362 | 0.513 | 0.80994 | 0.61128 | 1.3949 | 0.52683 | 0.85839 | 4.7943 | 0.45332 | 0.95586 | 1.6297 | 0.75476 | 0.36009 | 0.84715 | 0.47147 | 0.66046 | 0.66257 | 0.83766 | 0.54772
| Rel.Std.Dev. |  | 1.302% | 1.678% | 0.5751% | 1.249% | 1.912% | 2.605% | 3.411% | 1.629% | 1.382% | 1.35% | 2.1% | 1.824% | 5.117% | 1.273% | 2.173% | 11.36% | 1.051% | 2.53% | 4.211% | 1.869% | 0.9143% | 2.242% | 1.123% | 3.089% | 1.688% | 2.125% | 1.588%
| PDM3 |  | -1.093% | -0.6778% | 0.4442% | 3.062% | 3.603% | 0.3195% | 5.722% | 1.982% | 1.691% | -5.24% | -3.848% | -16.43% | -32.03% | 3.216% | -1.526% | 5.265% | 7.553% | -5.789% | -3.504% | 0.6936% | -1.8% | -5.789% | 4.668% | -46.68% | -2.132% | -1.717% | -13.98%
| Z-Score (Absolute) |  | 0 | 0.07 | 0.25 | 0.68 | 0.76 | 0.23 | 1.11 | 0.5 | 0.45 | 0.67 | 0.45 | 2.49 | 5.03 | 0.7 | 0.07 | 1.03 | 1.41 | 0.76 | 0.39 | 0.29 | 0.11 | 0.76 | 0.94 | 7.41 | 0.17 | 0.1 | 2.09
| Analytical results for P in OREAS 505b (Certified Value 0.086 wt.%) | [28 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 3
3
| 1 | 1 | 0.086 | 0.086 | 0.083 | 0.09 | 0.081 | 0.085 | 0.084 | 0.085 | 0.0869 | 0.085576 | 0.0925 | 0.09 | 0.09 | 0.085306 | 0.0889 | 0.085 | 0.088 | 0.0749 | 0.095 | 0.094991 | 0.084552 | 0.087 | 0.0791 | 0.071 |  |  |  | 1
| 1 | 2 | 0.086 | 0.083 | 0.086 | 0.09 | 0.085 | 0.087 | 0.084 | 0.086 | 0.0855 | 0.086405 | 0.0949 | 0.09 | 0.08 | 0.083993 | 0.0873 | 0.085 | 0.092 | 0.0745 | 0.093 | 0.091936 | 0.081657 | 0.087 | 0.0818 | 0.073 |  |  |  | 18
| 1 | 3 | 0.085 | 0.085 | 0.086 | 0.08 | 0.086 | 0.086 | 0.083 | 0.085 | 0.0866 | 0.085124 | 0.0919 | 0.09 | 0.09 | 0.083869 | 0.0868 | 0.084 | 0.091 | 0.0754 | 0.088 | 0.093267 | 0.083554 | 0.091 | 0.0788 | 0.072 |  |  |  | 16
| 1 | 4 | 0.087 | 0.087 | 0.084 | 0.08 | 0.084 | 0.086 | 0.086 | 0.086 | 0.087 | 0.085147 | 0.0968 | 0.09 | 0.09 | 0.083976 | 0.0876 | 0.086 | 0.092 | 0.0733 | 0.091 | 0.095359 | 0.083553 | 0.087 | 0.0838 | 0.072 |  |  |  | 0.085849
| 1 | 5 | 0.086 | 0.084 | 0.085 | 0.08 | 0.087 | 0.086 | 0.082 | 0.084 | 0.0864 | 0.086028 | 0.0938 | 0.09 | 0.08 | 0.085693 | 0.0877 | 0.086 | 0.093 | 0.0755 | 0.091 | 0.095559 | 0.082156 | 0.087 | 0.0813 | 0.07 |  |  |  | 44
| 1 | 6 | 0.085 | 0.087 | 0.084 | 0.08 | 0.084 | 0.086 | 0.084 | 0.084 | 0.0869 | 0.085857 | 0.095 | 0.09 | 0.09 | 0.085313 | 0.0881 | 0.083 | 0.09 | 0.076 | 0.088 | 0.095257 | 0.083354 | 0.087 | 0.0763 | 0.069
| Mean |  | 0.085833 | 0.085333 | 0.084667 | 0.083333 | 0.0845 | 0.086 | 0.083833 | 0.085 | 0.08655 | 0.08569 | 0.09415 | 0.09 | 0.086667 | 0.084692 | 0.087733 | 0.084833 | 0.091 | 0.074933 | 0.091 | 0.094395 | 0.083138 | 0.087667 | 0.080183 | 0.071167
| Median |  | 0.086 | 0.0855 | 0.0845 | 0.08 | 0.0845 | 0.086 | 0.084 | 0.085 | 0.08675 | 0.085716 | 0.09435 | 0.09 | 0.09 | 0.08465 | 0.08765 | 0.085 | 0.0915 | 0.07515 | 0.091 | 0.095124 | 0.083454 | 0.087 | 0.0802 | 0.0715
| Std Dev. |  | 0.00075277 | 0.001633 | 0.0012111 | 0.005164 | 0.0020736 | 0.00063246 | 0.0013292 | 0.00089443 | 0.00056125 | 0.0005062 | 0.0018008 | 1.5202e-17 | 0.005164 | 0.00083001 | 0.0007174 | 0.001169 | 0.0017889 | 0.00095219 | 0.0027568 | 0.001463 | 0.0010535 | 0.001633 | 0.0026498 | 0.001472
| Rel.Std.Dev. |  | 0.877% | 1.914% | 1.43% | 6.197% | 2.454% | 0.7354% | 1.585% | 1.052% | 0.6485% | 0.5907% | 1.913% | 1.689e-14% | 5.958% | 0.98% | 0.8177% | 1.378% | 1.966% | 1.271% | 3.029% | 1.55% | 1.267% | 1.863% | 3.305% | 2.068%
| PDM3 |  | -0.01859% | -0.601% | -1.378% | -2.931% | -1.572% | 0.1755% | -2.348% | -0.9893% | 0.8162% | -0.186% | 9.669% | 4.835% | 0.9521% | -1.348% | 2.195% | -1.183% | 6% | -12.72% | 6% | 9.954% | -3.158% | 2.117% | -6.6% | -17.1%
| Z-Score (Absolute) |  | 0.11 | 0.06 | 0.3 | 0.77 | 0.36 | 0.17 | 0.59 | 0.18 | 0.37 | 0.06 | 3.04 | 1.58 | 0.41 | 0.29 | 0.78 | 0.24 | 1.93 | 3.72 | 1.93 | 3.13 | 0.84 | 0.76 | 1.87 | 5.05
| Analytical results for Pb in OREAS 505b (Certified Value 49.7 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 1
2
| 1 | 1 | 51 | 43.5 | 48.12 | 50.8 | 49.3 | 45.4 | 52.3 | 50.4 | 52.3 | 43.7035 | 57.4146 | 45.9 | 53.22 | 33.95 | 51 | 52.1129 | 51 | 50.7 | 49.9 | 46.6 | 49.88 | 76.822 | 56.437 | 48.8 | 55.3 | 46 |  | 1
| 1 | 2 | 50 | 42.2 | 47.23 | 50.5 | 50 | 46.2 | 53.9 | 50.9 | 51.1 | 43.9677 | 54.3992 | 45 | 47.59 | 35.56 | 50 | 51.0196 | 50 | 53.8 | 49.9 | 48 | 50.18 | 47.441 | 52.426 | 47 | 50.3 | 54 |  | 19
| 1 | 3 | 50 | 43.2 | 48.46 | 52.1 | 49.6 | 45.9 | 55 | 50.7 | 51.1 | 44.8872 | 56.2636 | 45.4 | 54.36 | 35.64 | 51 | 50.0668 | 50 | 50.8 | 48.5 | 47.4 | 49.47 | 49.214 | 53.663 | 49.5 | 48.4 | 47 |  | 16
| 1 | 4 | 51 | 43.9 | 46.66 | 49.6 | 50.9 | 46.2 | 55.5 | 51.1 | 52.9 | 43.3303 | 56.8723 | 45.5 | 51.11 | 32.97 | 51 | 51.2307 | 50 | 51.1 | 51.5 | 47.8 | 49.57 | 49.688 | 53.591 | 46.7 | 52.3 | 47 |  | 49.7165
| 1 | 5 | 52 | 42.9 | 47.24 | 51.4 | 50.3 | 46.6 | 52.2 | 49.3 | 54.4 | 44.2955 | 56.6244 | 45.5 | 49.98 | 33.5 | 51 | 50.4214 | 50 | 51.5 | 50.1 | 47.1 | 50.18 | 49.503 | 53.14 | 48.3 | 51.7 | 46 |  | 45
| 1 | 6 | 52 | 44.2 | 48.01 | 53.7 | 50 | 45.3 | 53.8 | 50.1 | 51.8 | 44.4477 | 57.9792 | 44.8 | 51.92 | 32.74 | 51 | 51.8827 | 51 | 49.1 | 48.7 | 48.5 | 50.59 | 48.22 | 56.106 | 49.1 | 50.6 | 48
| Mean |  | 51 | 43.3167 | 47.62 | 51.35 | 50.0167 | 45.9333 | 53.7833 | 50.4167 | 52.2667 | 44.1053 | 56.5922 | 45.35 | 51.3633 | 34.06 | 50.8333 | 51.1224 | 50.3333 | 51.1667 | 49.7667 | 47.5667 | 49.9783 | 53.4813 | 54.2272 | 48.2333 | 51.4333 | 48
| Median |  | 51 | 43.35 | 47.625 | 51.1 | 50 | 46.05 | 53.85 | 50.55 | 52.05 | 44.1316 | 56.7483 | 45.45 | 51.515 | 33.725 | 51 | 51.1252 | 50 | 50.95 | 49.9 | 47.6 | 50.03 | 49.3585 | 53.627 | 48.55 | 51.15 | 47
| Std Dev. |  | 0.89443 | 0.71949 | 0.68203 | 1.4265 | 0.55648 | 0.50465 | 1.3527 | 0.65243 | 1.2565 | 0.55602 | 1.2326 | 0.3937 | 2.4064 | 1.2651 | 0.40825 | 0.79838 | 0.5164 | 1.528 | 1.0857 | 0.67725 | 0.42197 | 11.4663 | 1.647 | 1.1448 | 2.3218 | 3.0332
| Rel.Std.Dev. |  | 1.754% | 1.661% | 1.432% | 2.778% | 1.113% | 1.099% | 2.515% | 1.294% | 2.404% | 1.261% | 2.178% | 0.8681% | 4.685% | 3.714% | 0.8031% | 1.562% | 1.026% | 2.986% | 2.182% | 1.424% | 0.8443% | 21.44% | 3.037% | 2.374% | 4.514% | 6.319%
| PDM3 |  | 2.582% | -12.87% | -4.217% | 3.286% | 0.6037% | -7.609% | 8.18% | 1.408% | 5.129% | -11.29% | 13.83% | -8.783% | 3.312% | -31.49% | 2.246% | 2.828% | 1.241% | 2.917% | 0.1009% | -4.324% | 0.5266% | 7.573% | 9.073% | -2.983% | 3.453% | -3.453%
| Z-Score (Absolute) |  | 0.21 | 2.36 | 0.92 | 0.33 | 0.12 | 1.49 | 1.14 | 0.01 | 0.63 | 2.1 | 2.08 | 1.68 | 0.33 | 5.46 | 0.15 | 0.25 | 0.01 | 0.26 | 0.2 | 0.94 | 0.13 | 1.04 | 1.29 | 0.72 | 0.35 | 0.79
| Analytical results for Pr in OREAS 505b (Certified Value 7.8 ppm) | [28 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS |  | 2
3
| 1 | 1 | 1.01: 7.4; 1.02: 8.27; 1.03: 6.7219; 1.04: 8.3653; 1.05: 8.18; 1.06: 7.6; 1.07: 5.431 |  | 1
| 1 | 2 | 1.01: 7.3; 1.02: 8.31; 1.03: 7.2932; 1.04: 8.353; 1.05: 8.23; 1.06: 8.1; 1.07: 5.285 |  | 9
| 1 | 3 | 1.01: 7.5; 1.02: 8.22; 1.03: 6.9197; 1.04: 8.5496; 1.05: 7.83; 1.06: 7.7; 1.07: 5.458 |  | 16
| 1 | 4 | 1.01: 7.3; 1.02: 8.16; 1.03: 6.9181; 1.04: 8.5585; 1.05: 8.03; 1.06: 7.7; 1.07: 5.459 |  | 7.7969
| 1 | 5 | 1.01: 7.4; 1.02: 8.52; 1.03: 7.3943; 1.04: 8.2915; 1.05: 7.99; 1.06: 7.3; 1.07: 5.264 |  | 46
| 1 | 6 | 1.01: 7.3; 1.02: 8.8; 1.03: 7.1075; 1.04: 8.2494; 1.05: 7.83; 1.06: 7.5; 1.07: 5.585
| Mean |  | 1.01: 7.3667; 1.02: 8.38; 1.03: 7.0591; 1.04: 8.3946; 1.05: 8.015; 1.06: 7.65; 1.07: 5.4137
| Median |  | 1.01: 7.35; 1.02: 8.29; 1.03: 7.0136; 1.04: 8.3592; 1.05: 8.01; 1.06: 7.65; 1.07: 5.4445
| Std Dev. |  | 1.01: 0.08165; 1.02: 0.23958; 1.03: 0.25398; 1.04: 0.13053; 1.05: 0.16897; 1.06: 0.26646; 1.07: 0.12052
| Rel.Std.Dev. |  | 1.01: 1.108%; 1.02: 2.859%; 1.03: 3.598%; 1.04: 1.555%; 1.05: 2.108%; 1.06: 3.483%; 1.07: 2.226%
| PDM3 |  | 1.01: -5.518%; 1.02: 7.479%; 1.03: -9.462%; 1.04: 7.666%; 1.05: 2.797%; 1.06: -1.884%; 1.07: -30.57%
| Z-Score (Absolute) |  | 1.01: 0.32; 1.02: 0.83; 1.03: 0.67; 1.04: 0.85; 1.05: 0.42; 1.06: 0; 1.07: 2.55
| Analytical results for Rb in OREAS 505b (Certified Value 153 ppm) | [28 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 0
0
| 1 | 1 | 152 | 166.9 | 136.2 | 159 | 152.5 | 150 | 145 | 147 | 160.16 | 169.1389 | 158 | 163 | 163.3515 | 150 | 149 | 151 | 122.8 | 125.345 | 125.1 | 155.5 | [7 empty] | 1
| 1 | 2 | 155 | 165.7 | 136.4 | 156 | 149 | 151.5 | 138 | 151.5 | 158.23 | 167.2723 | 154 | 156 | 162.6031 | 149 | 127.5 | 152 | 147.5 | 121.383 | 132.6 | 157.4 | [7 empty] | 37
| 1 | 3 | 154 | 151.9 | 140.2 | 155 | 154 | 150.5 | 166 | 152.5 | 162.03 | 167.8547 | 146 | 158 | 161.3553 | 145.7 | 137 | 148 | 147.5 | 125.371 | 143.9 | 157.1 | [7 empty] | 16
| 1 | 4 | 157 | 163.8 | 137.5 | 164 | 154.5 | 150.5 | 168.5 | 149 | 163.66 | 168.9924 | 159 | 171 | 163.1275 | 146.8 | 151 | 153.5 | 149.5 | 124.252 | 140.4 | 158.5 | [7 empty] | 152.7374
| 1 | 5 | 155 | 169 | 138 | 155 | 154.5 | 150 | 156 | 143.5 | 161.11 | 163.3872 | 156 | 160 | 161.7261 | 146 | 144.5 | 155 | 128.7 | 122.076 | 137.6 | 154.1 | [7 empty] | 47
| 1 | 6 | 156 | 130.8 | 138.3 | 155 | 154 | 150 | 162 | 152 | 162 | 166.2015 | 149 | 168 | 163.1001 | 145.5 | 146 | 149 | 116.8 | 125.575 | 139.6 | 156
| Mean |  | 154.8333 | 158.0167 | 137.7667 | 157.3333 | 153.0833 | 150.4167 | 155.9167 | 149.25 | 161.1983 | 167.1412 | 153.6667 | 162.6667 | 162.5439 | 147.1667 | 142.5 | 151.4167 | 135.4667 | 124.0003 | 136.5333 | 156.4333
| Median |  | 155 | 164.75 | 137.75 | 155.5 | 154 | 150.25 | 159 | 150.25 | 161.555 | 167.5635 | 155 | 161.5 | 162.8516 | 146.4 | 145.25 | 151.5 | 138.1 | 124.7985 | 138.6 | 156.55
| Std Dev. |  | 1.7224 | 14.63 | 1.4597 | 3.6148 | 2.1311 | 0.58452 | 12.1426 | 3.5036 | 1.8596 | 2.141 | 5.164 | 5.8538 | 0.82307 | 1.8875 | 8.7864 | 2.6536 | 14.4306 | 1.8321 | 6.7224 | 1.5591
| Rel.Std.Dev. |  | 1.112% | 9.259% | 1.06% | 2.298% | 1.392% | 0.3886% | 7.788% | 2.347% | 1.154% | 1.281% | 3.361% | 3.599% | 0.5064% | 1.283% | 6.166% | 1.753% | 10.65% | 1.478% | 4.924% | 0.9966%
| PDM3 |  | 1.372% | 3.456% | -9.802% | 3.009% | 0.2265% | -1.519% | 2.082% | -2.283% | 5.54% | 9.43% | 0.6084% | 6.501% | 6.421% | -3.647% | -6.703% | -0.8647% | -11.31% | -18.81% | -10.61% | 2.42%
| Z-Score (Absolute) |  | 0.18 | 0.58 | 1.94 | 0.49 | 0.04 | 0.37 | 0.32 | 0.51 | 0.97 | 1.71 | 0.04 | 1.15 | 1.14 | 0.77 | 1.35 | 0.24 | 2.23 | 3.65 | 2.09 | 0.38
| Analytical results for Re in OREAS 505b (Certified Value 0.023 ppm) | [28 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.12 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [11 empty] | 3
3
| 1 | 1 | < 0.05 | 0.023 | 0.02 | 0.022 | 0.025 | 0.023 | 0.021 | 0.019 | 0.019 | < 0.05 | 0.023 | 0.022 | 0.023 | 0.018 | 0.032 | 0.023 | [11 empty] | 1
| 1 | 2 | < 0.05 | 0.022 | 0.02 | 0.022 | 0.024 | 0.024 | 0.022 | 0.019 | 0.021 | < 0.05 | 0.022 | 0.022 | 0.023 | 0.027 | 0.028 | 0.021 | [11 empty] | 38
| 1 | 3 | < 0.05 | 0.02 | 0.023 | 0.022 | 0.025 | 0.026 | 0.026 | 0.024 | 0.023 | < 0.05 | 0.026 | 0.023 | 0.021 | 0.018 | 0.03 | 0.025 | [11 empty] | 16
| 1 | 4 | < 0.05 | 0.02 | 0.019 | 0.024 | 0.023 | 0.026 | 0.026 | 0.019 | 0.023 | < 0.05 | 0.025 | 0.022 | 0.024 | 0.027 | 0.032 | 0.026 | [11 empty] | 0.022515
| 1 | 5 | < 0.05 | 0.021 | 0.021 | 0.025 | 0.022 | 0.023 | 0.024 | 0.022 | 0.019 | < 0.05 | 0.021 | 0.024 | 0.023 | 0.018 | 0.029 | 0.023 | [11 empty] | 48
| 1 | 6 | < 0.05 | 0.015 | 0.021 | 0.024 | 0.019 | 0.026 | 0.024 | 0.026 | 0.025 | < 0.05 | 0.023 | 0.021 | 0.023 | 0.018 | 0.029 | 0.023
| Mean |  |  | 0.020167 | 0.020667 | 0.023167 | 0.023 | 0.024667 | 0.023833 | 0.0215 | 0.021667 |  | 0.023333 | 0.022333 | 0.022833 | 0.021 | 0.03 | 0.0235
| Median |  |  | 0.0205 | 0.0205 | 0.023 | 0.0235 | 0.025 | 0.024 | 0.0205 | 0.022 |  | 0.023 | 0.022 | 0.023 | 0.018 | 0.0295 | 0.023
| Std Dev. |  |  | 0.0027869 | 0.0013663 | 0.0013292 | 0.0022804 | 0.0015055 | 0.0020412 | 0.0030166 | 0.0024221 |  | 0.0018619 | 0.0010328 | 0.00098319 | 0.0046476 | 0.0016733 | 0.0017607
| Rel.Std.Dev. |  |  | 13.82% | 6.611% | 5.737% | 9.915% | 6.104% | 8.565% | 14.03% | 11.18% |  | 7.98% | 4.624% | 4.306% | 22.13% | 5.578% | 7.492%
| PDM3 |  |  | -10.43% | -8.211% | 2.893% | 2.152% | 9.555% | 5.854% | -4.51% | -3.77% |  | 3.633% | -0.8086% | 1.412% | -6.73% | 33.24% | 4.373%
| Z-Score (Absolute) |  | 0.86 | 1.31 | 1.09 | 0.04 | 0.04 | 0.71 | 0.34 | 0.71 | 0.64 | 0.86 | 0.11 | 0.34 | 0.11 | 0.94 | 3.11 | 0.19
| Analytical results for S in OREAS 505b (Certified Value 0.537 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 0.529 | 0.6 | 0.49 | 0.56 | 0.51 | 0.55 | 0.56 | 0.54 | 0.534 | 0.49077 | 0.57003 | 0.56 | 0.51 | 0.55 | 0.53379 | 0.53 | 0.53 | 0.57 | 0.55 | 0.53 | 0.533 | 0.51424 | 0.49 | 0.53 | 0.544 |  |  | 1
| 1 | 2 | 0.531 | 0.5 | 0.49 | 0.55 | 0.54 | 0.57 | 0.56 | 0.55 | 0.5376 | 0.50167 | 0.57806 | 0.58 | 0.5 | 0.53 | 0.53484 | 0.52 | 0.54 | 0.59 | 0.55 | 0.54 | 0.514 | 0.51504 | 0.49 | 0.55 | 0.561 |  |  | 23
| 1 | 3 | 0.53 | 0.6 | 0.49 | 0.56 | 0.55 | 0.55 | 0.55 | 0.53 | 0.5523 | 0.4937 | 0.56331 | 0.56 | 0.5 | 0.53 | 0.52419 | 0.54 | 0.54 | 0.58 | 0.54 | 0.51 | 0.549 | 0.51535 | 0.51 | 0.52 | 0.545 |  |  | 16
| 1 | 4 | 0.534 | 0.6 | 0.49 | 0.54 | 0.53 | 0.56 | 0.57 | 0.55 | 0.5486 | 0.49083 | 0.56972 | 0.56 | 0.5 | 0.55 | 0.52206 | 0.54 | 0.54 | 0.59 | 0.55 | 0.52 | 0.602 | 0.51651 | 0.49 | 0.53 | 0.545 |  |  | 0.53693
| 1 | 5 | 0.53 | 0.5 | 0.5 | 0.55 | 0.55 | 0.56 | 0.55 | 0.54 | 0.552 | 0.49715 | 0.56584 | 0.57 | 0.51 | 0.54 | 0.52672 | 0.52 | 0.54 | 0.6 | 0.55 | 0.54 | 0.545 | 0.51726 | 0.5 | 0.54 | 0.54 |  |  | 49
| 1 | 6 | 0.539 | 0.6 | 0.49 | 0.54 | 0.53 | 0.55 | 0.56 | 0.52 | 0.5508 | 0.48017 | 0.57003 | 0.57 | 0.5 | 0.55 | 0.53313 | 0.55 | 0.53 | 0.57 | 0.57 | 0.53 | 0.558 | 0.51602 | 0.5 | 0.55 | 0.541
| Mean |  | 0.53217 | 0.56667 | 0.49167 | 0.55 | 0.535 | 0.55667 | 0.55833 | 0.53833 | 0.54588 | 0.49238 | 0.5695 | 0.56667 | 0.50333 | 0.54167 | 0.52912 | 0.53333 | 0.53667 | 0.58333 | 0.55167 | 0.52833 | 0.55017 | 0.51574 | 0.49667 | 0.53667 | 0.546
| Median |  | 0.5305 | 0.6 | 0.49 | 0.55 | 0.535 | 0.555 | 0.56 | 0.54 | 0.5497 | 0.49226 | 0.56987 | 0.565 | 0.5 | 0.545 | 0.52993 | 0.535 | 0.54 | 0.585 | 0.55 | 0.53 | 0.547 | 0.51569 | 0.495 | 0.535 | 0.5445
| Std Dev. |  | 0.0037639 | 0.05164 | 0.0040825 | 0.0089443 | 0.015166 | 0.008165 | 0.0075277 | 0.01169 | 0.0079996 | 0.0072787 | 0.0050122 | 0.008165 | 0.005164 | 0.0098319 | 0.0054883 | 0.012111 | 0.005164 | 0.012111 | 0.0098319 | 0.01169 | 0.029593 | 0.0010846 | 0.008165 | 0.012111 | 0.007642
| Rel.Std.Dev. |  | 0.7073% | 9.113% | 0.8303% | 1.626% | 2.835% | 1.467% | 1.348% | 2.172% | 1.465% | 1.478% | 0.8801% | 1.441% | 1.026% | 1.815% | 1.037% | 2.271% | 0.9622% | 2.076% | 1.782% | 2.213% | 5.379% | 0.2103% | 1.644% | 2.257% | 1.4%
| PDM3 |  | -0.8869% | 5.539% | -8.43% | 2.434% | -0.3592% | 3.676% | 3.986% | 0.2616% | 1.668% | -8.297% | 6.066% | 5.539% | -6.257% | 0.8824% | -1.454% | -0.6696% | -0.0488% | 8.643% | 2.745% | -1.601% | 2.465% | -3.947% | -7.499% | -0.0488% | 1.689%
| Z-Score (Absolute) |  | 0.29 | NA | 2.46 | 0.67 | 0.13 | 1.03 | 1.12 | 0.04 | 0.45 | 2.42 | 1.71 | 1.56 | 1.83 | 0.22 | 0.45 | 0.22 | 0.04 | 2.46 | 0.76 | 0.49 | 0.68 | 1.17 | 2.19 | 0.04 | 0.46
Results from laboratory 1.02 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Sb in OREAS 505b (Certified Value 10.1 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  | 1
2
| 1 | 1 | 9.8 | 9.5 | 8.92 | 9.86 | 9.9 | 10.1 | 10.1 | 9.89 | 10.53 | 9.6341 | 12.6 | 5.91 | 6.83 | 10.3 | 10.6027 | 9.46 | 10.4 | 9.33 | 11 | 10.51 | < 10 | 10.166 | 10.1 | 9.8 | 13 |  |  | 1
| 1 | 2 | 10.2 | 9.9 | 9.12 | 10.04 | 10.15 | 10.4 | 10.4 | 9.69 | 10.47 | 9.6952 | 12.2 | 4.26 | 7.09 | 9.55 | 10.5678 | 9.62 | 9.18 | 9.31 | 10 | 10.71 | < 10 | 10.126 | 10.4 | 10.3 | 9 |  |  | 39
| 1 | 3 | 10.3 | 9.7 | 8.98 | 10.03 | 10.25 | 10.2 | 10.85 | 10.25 | 10.31 | 9.8694 | 11.8 | 6.45 | 6.7 | 9.64 | 10.569 | 9.36 | 9.84 | 9.14 | 10 | 10.4 | < 10 | 10.328 | 10.4 | 9.2 | 10 |  |  | 16
| 1 | 4 | 10.3 | 9.6 | 8.95 | 10.3 | 10.4 | 10.3 | 10.7 | 9.51 | 10.83 | 9.8558 | 12.7 | 8.5 | 7.13 | 10.6 | 10.6584 | 9.61 | 10.55 | 9.33 | 10 | 10.4 | < 10 | 10.177 | 10.2 | 9.93 | 10 |  |  | 10.0879
| 1 | 5 | 10.3 | 10.4 | 9.03 | 10.05 | 10.75 | 10.25 | 10.15 | 9.67 | 11.07 | 9.844 | 12.1 | 4.92 | 6.8 | 9.51 | 10.5051 | 9.5 | 9.93 | 9.6 | 11 | 10.3 | < 10 | 10.253 | 10.4 | 9.96 | 7 |  |  | 50
| 1 | 6 | 10.1 | 9.9 | 9.01 | 10.11 | 10.3 | 9.68 | 10.4 | 10.05 | 10.73 | 9.8937 | 13 | 5.83 | 6.63 | 9.83 | 10.6937 | 9.46 | 10.25 | 9.33 | 11 | 10.81 | < 10 | 10.426 | 10.1 | 9.4 | 11
| Mean |  | 10.1667 | 9.8333 | 9.0017 | 10.065 | 10.2917 | 10.155 | 10.4333 | 9.8433 | 10.6567 | 9.7987 | 12.4 | 5.9783 | 6.8633 | 9.905 | 10.5995 | 9.5017 | 10.025 | 9.34 | 10.5 | 10.5217 |  | 10.246 | 10.2667 | 9.765 | 10
| Median |  | 10.25 | 9.8 | 8.995 | 10.045 | 10.275 | 10.225 | 10.4 | 9.79 | 10.63 | 9.8499 | 12.4 | 5.87 | 6.815 | 9.735 | 10.5858 | 9.48 | 10.09 | 9.33 | 10.5 | 10.455 |  | 10.215 | 10.3 | 9.865 | 10
| Std Dev. |  | 0.19664 | 0.32042 | 0.070261 | 0.14237 | 0.28181 | 0.25328 | 0.29609 | 0.27384 | 0.27471 | 0.10691 | 0.44272 | 1.4616 | 0.20432 | 0.44653 | 0.068132 | 0.099281 | 0.4948 | 0.14751 | 0.54772 | 0.19874 |  | 0.11378 | 0.15055 | 0.40129 | 2
| Rel.Std.Dev. |  | 1.934% | 3.258% | 0.7805% | 1.415% | 2.738% | 2.494% | 2.838% | 2.782% | 2.578% | 1.091% | 3.57% | 24.45% | 2.977% | 4.508% | 0.6428% | 1.045% | 4.936% | 1.579% | 5.216% | 1.889% |  | 1.111% | 1.466% | 4.109% | 20%
| PDM3 |  | 0.7806% | -2.524% | -10.77% | -0.2272% | 2.02% | 0.665% | 3.424% | -2.425% | 5.638% | -2.867% | 22.92% | -40.74% | -31.96% | -1.813% | 5.071% | -5.811% | -0.6237% | -7.414% | 4.085% | 4.3% |  | 1.567% | 1.772% | -3.201% | -0.8715%
| Z-Score (Absolute) |  | 0.36 | 0.48 | 2.59 | 0.1 | 0.67 | 0.33 | 1.03 | 0.46 | 1.6 | 0.57 | 6.01 | 10.23 | 7.99 | 0.3 | 1.45 | 1.32 | 0 | 1.73 | NA | 1.26 | 12.71 | 0.56 | 0.61 | 0.66 | NA
Results from laboratories 1.22 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 505b (Certified Value 8.58 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.26 | 1.27 | 1.29 | 1.30 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  |  | 2
3
| 1 | 1 | 8 | 8 | 7.5 | 9.1 | 8.4 | 8.1 | 8.8 | 8.3 | 8.9 | 8.4048 | 8.4823 | 8.8 | 7.54 | 9.6 | 9.4753 | 8.8 | 8.9 | 8.7 | 8.2 | 8.1947 | 8.1 | 8.3 | 7.8 |  |  |  |  | 1
| 1 | 2 | 8 | 8 | 7.6 | 9.1 | 8.3 | 8.3 | 9.2 | 8.2 | 8.9 | 8.8997 | 8.3314 | 9 | 8.03 | 8.9 | 9.1474 | 8.8 | 8.8 | 8.9 | 8.1 | 7.966 | 8.3 | 8.3 | 8.3 |  |  |  |  | 40
| 1 | 3 | 8 | 8 | 7.6 | 9.2 | 8.5 | 8.3 | 9.5 | 8.8 | 9.1 | 8.9242 | 8.2857 | 8.9 | 8.71 | 8.6 | 9.2302 | 8.4 | 8.5 | 8.5 | 8.1 | 8.1351 | 8.7 | 8.3 | 8.3 |  |  |  |  | 16
| 1 | 4 | 8 | 8 | 7.5 | 9.2 | 8.7 | 8.2 | 9.5 | 8.2 | 9.2 | 8.761 | 8.3754 | 8.9 | 8.82 | 9.6 | 9.2692 | 8.7 | 8.8 | 9.1 | 8.2 | 8.1444 | 8.4 | 8.4 | 8.3 |  |  |  |  | 8.578
| 1 | 5 | 8 | 8 | 7.7 | 8.8 | 8.8 | 8.1 | 8.9 | 8.2 | 9.4 | 9.3805 | 8.3904 | 8.9 | 8.14 | 9 | 9.2824 | 8.6 | 8.8 | 9 | 8 | 7.9956 | 7.7 | 8.4 | 7.9 |  |  |  |  | 51
| 1 | 6 | 8 | 8 | 7.6 | 8.7 | 8.6 | 7.9 | 9.2 | 8.7 | 9.1 | 9.3 | 8.451 | 9.2 | 8.61 | 9.7 | 9.4119 | 8.6 | 8.5 | 8.7 | 8.2 | 8.1049 | 8.3 | 7.7 | 7.9
| Mean |  | 8 | 8 | 7.5833 | 9.0167 | 8.55 | 8.15 | 9.1833 | 8.4 | 9.1 | 8.945 | 8.386 | 8.95 | 8.3083 | 9.2333 | 9.3028 | 8.65 | 8.7167 | 8.8167 | 8.1333 | 8.0901 | 8.25 | 8.2333 | 8.0833
| Median |  | 8 | 8 | 7.6 | 9.1 | 8.55 | 8.15 | 9.2 | 8.25 | 9.1 | 8.9119 | 8.3829 | 8.9 | 8.375 | 9.3 | 9.2758 | 8.65 | 8.8 | 8.8 | 8.15 | 8.12 | 8.3 | 8.3 | 8.1
| Std Dev. |  | 0 | 0 | 0.075277 | 0.2137 | 0.18708 | 0.15166 | 0.29269 | 0.27568 | 0.18974 | 0.35884 | 0.073029 | 0.13784 | 0.49199 | 0.45898 | 0.12052 | 0.15166 | 0.17224 | 0.22286 | 0.08165 | 0.089958 | 0.33317 | 0.26583 | 0.24014
| Rel.Std.Dev. |  | 0% | 0% | 0.9927% | 2.37% | 2.188% | 1.861% | 3.187% | 3.282% | 2.085% | 4.012% | 0.8708% | 1.54% | 5.922% | 4.971% | 1.295% | 1.753% | 1.976% | 2.528% | 1.004% | 1.112% | 4.038% | 3.229% | 2.971%
| PDM3 |  | -6.739% | -6.739% | -11.6% | 5.113% | -0.3269% | -4.99% | 7.056% | -2.076% | 6.085% | 4.278% | -2.238% | 4.336% | -3.144% | 7.639% | 8.449% | 0.8388% | 1.616% | 2.782% | -5.184% | -5.688% | -3.824% | -4.019% | -5.767%
| Z-Score (Absolute) |  | NA | NA | 1.63 | 0.79 | 0 | 0.67 | 1.07 | 0.25 | 0.93 | 0.67 | 0.28 | 0.67 | 0.41 | 1.15 | 1.27 | 0.17 | 0.28 | 0.45 | 0.7 | 0.78 | 0.51 | 0.53 | 0.79
Results from laboratories 1.01 and 1.02 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 505b (Certified Value 3.98 ppm) | [28 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.29 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 4 | 5 | 3.4 | 4 | 4 | 5 | 4 | 3 | 3.8 | 3.0764 | 13.4 | 2.86 | 3 | 4.1974 | 4 | 4 | 6 | < 5 | 4.5 | < 10 | 1.3808 | 4 | [5 empty] | 1
| 1 | 2 | 4 | 5 | 3.5 | 4 | 3 | 5 | 5 | 3 | 4 | 3.7066 | 11.8 | 0.59 | 4 | 3.8584 | 4 | 4 | 6 | < 5 | 3.6 | < 10 | 1.6143 | 4 | [5 empty] | 41
| 1 | 3 | 4 | 6 | 3.7 | 4 | 4 | 5 | 5 | 5 | 4.2 | 3.6114 | 13.9 | 2.44 | 4 | 3.0661 | 4 | 4 | 5 | < 5 | 4.1 | < 10 | 0.90212 | 5 | [5 empty] | 16
| 1 | 4 | 4 | 4 | 3.8 | 4 | 3 | 5 | 4 | 4 | 4.4 | 3.4984 | 13.1 | 0.55 | 4 | 3.667 | 4 | 4 | 6 | < 5 | 4 | < 10 | 1.7259 | 4 | [5 empty] | 3.9828
| 1 | 5 | 4 | 6 | 3.6 | 4 | 3 | 5 | 4 | 3 | 4.1 | 3.2187 | 11.4 | 1.38 | 4 | 3.2816 | 4 | 4 | 6 | < 5 | 4 | < 10 | 1.5683 | 4 | [5 empty] | 52
| 1 | 6 | 4 | 5 | 3.4 | 4 | 3 | 5 | 4 | 4 | 4.4 | 3.572 | 10.1 | 0.24 | 4 | 2.7972 | 4 | 4 | 6 | < 5 | 3.8 | < 10 | 2.0699 | 4
| Mean |  | 4 | 5.1667 | 3.5667 | 4 | 3.3333 | 5 | 4.3333 | 3.6667 | 4.15 | 3.4473 | 12.2833 | 1.3433 | 3.8333 | 3.478 | 4 | 4 | 5.8333 |  | 4 |  | 1.5436 | 4.1667
| Median |  | 4 | 5 | 3.55 | 4 | 3 | 5 | 4 | 3.5 | 4.15 | 3.5352 | 12.45 | 0.985 | 4 | 3.4743 | 4 | 4 | 6 |  | 4 |  | 1.5913 | 4
| Std Dev. |  | 0 | 0.75277 | 0.1633 | 0 | 0.5164 | 0 | 0.5164 | 0.8165 | 0.23452 | 0.24581 | 1.4359 | 1.088 | 0.40825 | 0.52339 | 0 | 0 | 0.40825 |  | 0.30332 |  | 0.38829 | 0.40825
| Rel.Std.Dev. |  | 0% | 14.57% | 4.578% | 0% | 15.49% | 0% | 11.92% | 22.27% | 5.651% | 7.131% | 11.69% | 80.99% | 10.65% | 15.05% | 0% | 0% | 6.999% |  | 7.583% |  | 25.16% | 9.798%
| PDM3 |  | 0.4312% | 29.72% | -10.45% | 0.4312% | -16.31% | 25.54% | 8.8% | -7.938% | 4.197% | -13.45% | 208.4% | -66.27% | -3.753% | -12.68% | 0.4312% | 0.4312% | 46.46% |  | 0.4312% |  | -61.24% | 4.616%
| Z-Score (Absolute) |  | 0 | 1.65 | 0.61 | 0 | 0.94 | 1.41 | 0.47 | 0.47 | 0.21 | 0.78 | 11.69 | 3.75 | 0.24 | 0.74 | 0 | 0 | 2.59 | 2.12 | 0 | 1.41 | 3.47 | 0.24
| Analytical results for Si in OREAS 505b (Indicative Value 37.4 wt.%) | [28 empty] | Indicative
Si | Round | Replicate | 1.01: Lab | [5 empty] | 1
| No | No | 1.01: 1.13 | [5 empty] | wt.%
|  |  | 1.01: 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 1.01: 39.2438 | [5 empty] | 1
| 1 | 2 | 1.01: 32.6834 | [5 empty] | 4
| 1 | 3 | 1.01: 41.2828 | [5 empty] | 16
| 1 | 4 | 1.01: 37.3193 | [5 empty] | 37.4025
| 1 | 5 | 1.01: 32.8828 | [5 empty] | 10
| 1 | 6 | 1.01: 41.0028
| Mean |  | 1.01: 37.4025
| Median |  | 1.01: 38.2815
| Std Dev. |  | 1.01: 3.85
| Rel.Std.Dev. |  | 1.01: 10.29%
PDM3 | bar: 1.01
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sm in OREAS 505b (Certified Value 6.26 ppm) | [28 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [5 empty] | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 | [5 empty] | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.01: 5.9; 1.02: 6.64; 1.03: 5.0718; 1.04: 6.7083; 1.05: 6.73; 1.06: 6.1; 1.07: 1.3808 | [5 empty] | 1
| 1 | 2 | 1.01: 5.7; 1.02: 6.67; 1.03: 5.8292; 1.04: 6.7813; 1.05: 6.7; 1.06: 6.6; 1.07: 1.6143 | [5 empty] | 7
| 1 | 3 | 1.01: 5.9; 1.02: 6.76; 1.03: 5.3098; 1.04: 6.818; 1.05: 6.35; 1.06: 6.4; 1.07: 0.90212 | [5 empty] | 16
| 1 | 4 | 1.01: 5.8; 1.02: 6.59; 1.03: 5.2344; 1.04: 6.8481; 1.05: 6.46; 1.06: 6.4; 1.07: 1.7259 | [5 empty] | 6.2598
| 1 | 5 | 1.01: 5.8; 1.02: 6.91; 1.03: 5.4975; 1.04: 6.7424; 1.05: 6.38; 1.06: 6.1; 1.07: 1.5683 | [5 empty] | 53
| 1 | 6 | 1.01: 5.9; 1.02: 7.01; 1.03: 5.5424; 1.04: 6.6811; 1.05: 6.29; 1.06: 6.2; 1.07: 2.0699
| Mean |  | 1.01: 5.8333; 1.02: 6.7633; 1.03: 5.4142; 1.04: 6.7632; 1.05: 6.485; 1.06: 6.3; 1.07: 1.5436
| Median |  | 1.01: 5.85; 1.02: 6.715; 1.03: 5.4036; 1.04: 6.7619; 1.05: 6.42; 1.06: 6.3; 1.07: 1.5913
| Std Dev. |  | 1.01: 0.08165; 1.02: 0.16513; 1.03: 0.26686; 1.04: 0.064369; 1.05: 0.18663; 1.06: 0.2; 1.07: 0.38829
| Rel.Std.Dev. |  | 1.01: 1.4%; 1.02: 2.441%; 1.03: 4.929%; 1.04: 0.9518%; 1.05: 2.878%; 1.06: 3.175%; 1.07: 25.16%
| PDM3 |  | 1.01: -6.813%; 1.02: 8.043%; 1.03: -13.51%; 1.04: 8.041%; 1.05: 3.597%; 1.06: 0.6415%; 1.07: -75.34%
| Z-Score (Absolute) |  | 1.01: 0.68; 1.02: 0.67; 1.03: 1.29; 1.04: 0.67; 1.05: 0.27; 1.06: 0; 1.07: 6.92
| Analytical results for Sn in OREAS 505b (Certified Value 4.41 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 2
3
| 1 | 1 | 4.6 | 4.2 | 3.7 | 4.5 | 4.2 | 4.2 | 4.3 | 4.3 | 4.4 | 6.145 | 4.6683 | 5.2 | 8.16 | < 5 | 4.6 | 4.9053 | 4.1 | 4.4 | 3.9 | < 10 | 4.5 | < 10 | 5.9938 | 4.8 | 5 | < 10 |  | 1
| 1 | 2 | 4.5 | 4.1 | 4 | 4.5 | 4.2 | 4.2 | 4.4 | 4.3 | 4.3 | 4.1595 | 4.643 | 5.5 | 6.1 | < 5 | 4.3 | 4.9123 | 4.1 | 3.5 | 3.9 | < 10 | 4.6 | < 10 | 5.4833 | 4.7 | 4.5 | < 10 |  | 14
| 1 | 3 | 4.7 | 3.6 | 3.9 | 4.3 | 4.2 | 4.3 | 4.5 | 4.5 | 4.7 | 3.7373 | 4.7194 | 5.3 | 7.24 | < 5 | 4.3 | 4.8107 | 3.8 | 4 | 3.7 | < 10 | 4.4 | < 10 | 5.4428 | 4.7 | 4.3 | < 10 |  | 16
| 1 | 4 | 4.7 | 4 | 3.8 | 4.6 | 4.4 | 4.3 | 4.5 | 4.2 | 4.7 | 4.9298 | 4.9937 | 4.9 | 6.73 | < 5 | 4.5 | 4.8442 | 4 | 4.4 | 4 | < 10 | 4.4 | < 10 | 5.7469 | 4.6 | 4.4 | < 10 |  | 4.4083
| 1 | 5 | 4.7 | 4.1 | 3.9 | 4.3 | 4.4 | 4.2 | 4.3 | 4.2 | 4.8 | 5.1655 | 4.7588 | 5.4 | 12.07 | < 5 | 4.3 | 4.8023 | 3.8 | 4.1 | 4.3 | < 10 | 4.5 | < 10 | 5.4447 | 4.8 | 4.3 | < 10 |  | 54
| 1 | 6 | 4.7 | 4.1 | 3.8 | 4.3 | 4.2 | 4 | 4.4 | 4.6 | 4.7 | 4.7655 | 4.7122 | 5.1 | 6.76 | < 5 | 4.4 | 4.9663 | 3.8 | 4.2 | 4 | < 10 | 4.5 | < 10 | 5.747 | 4.9 | 4.2 | < 10
| Mean |  | 4.65 | 4.0167 | 3.85 | 4.4167 | 4.2667 | 4.2 | 4.4 | 4.35 | 4.6 | 4.8171 | 4.7492 | 5.2333 | 7.8433 |  | 4.4 | 4.8735 | 3.9333 | 4.1 | 3.9667 |  | 4.4833 |  | 5.6431 | 4.75 | 4.45
| Median |  | 4.7 | 4.1 | 3.85 | 4.4 | 4.2 | 4.2 | 4.4 | 4.3 | 4.7 | 4.8476 | 4.7158 | 5.25 | 7 |  | 4.35 | 4.8747 | 3.9 | 4.15 | 3.95 |  | 4.5 |  | 5.6151 | 4.75 | 4.35
| Std Dev. |  | 0.083666 | 0.2137 | 0.10488 | 0.13292 | 0.10328 | 0.10954 | 0.089443 | 0.16432 | 0.2 | 0.83648 | 0.12644 | 0.21602 | 2.1809 |  | 0.12649 | 0.06482 | 0.15055 | 0.33466 | 0.19664 |  | 0.075277 |  | 0.22342 | 0.10488 | 0.2881
| Rel.Std.Dev. |  | 1.799% | 5.32% | 2.724% | 3.009% | 2.421% | 2.608% | 2.033% | 3.777% | 4.348% | 17.36% | 2.662% | 4.128% | 27.81% |  | 2.875% | 1.33% | 3.828% | 8.163% | 4.957% |  | 1.679% |  | 3.959% | 2.208% | 6.474%
| PDM3 |  | 5.484% | -8.883% | -12.66% | 0.1905% | -3.212% | -4.725% | -0.1876% | -1.322% | 4.349% | 9.274% | 7.735% | 18.72% | 77.92% |  | -0.1876% | 10.55% | -10.77% | -6.993% | -10.02% |  | 1.703% |  | 28.01% | 7.752% | 0.9466%
| Z-Score (Absolute) |  | 0.34 | 0.85 | 1.16 | 0.09 | 0.38 | 0.5 | 0.13 | 0.22 | 0.25 | 0.66 | 0.53 | 1.44 | 6.35 | 3.7 | 0.13 | 0.77 | 1 | 0.69 | 0.94 | 1 | 0.03 | 1 | 2.21 | 0.53 | 0.03 | 1
| Analytical results for Sr in OREAS 505b (Certified Value 266 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 275 | 269 | 245 | 278 | 256 | 262 | 274 | 266 | 271.93 | 268.4807 | 264 | 282.73 | 290.92 | 265.4 | 256 | 283.8881 | 265 | 264 | 264 | 239 | 283 | 239.8 | 295.0756 | 269 | 248.6 | 259 |  | 1
| 1 | 2 | 279 | 255 | 248 | 272 | 266 | 273 | 280 | 269 | 268.75 | 267.552 | 274 | 272.89 | 286.53 | 264.8 | 251 | 285.6253 | 259 | 268 | 277 | 237 | 301 | 243.16 | 288.4298 | 270 | 253.3 | 261 |  | 15
| 1 | 3 | 275 | 256 | 249 | 278 | 266 | 264 | 276 | 269 | 272.67 | 263.5653 | 263 | 276.86 | 293.23 | 263.2 | 257 | 279.2226 | 262 | 261 | 272 | 236 | 267 | 227.55 | 293.8012 | 282 | 248.1 | 259 |  | 16
| 1 | 4 | 282 | 275 | 243 | 268 | 263 | 268 | 288 | 269 | 278.87 | 259.4137 | 266 | 281.67 | 288.23 | 261.4 | 264 | 284.2964 | 265 | 265 | 278 | 251 | 280 | 244.03 | 295.683 | 270 | 251.1 | 260 |  | 265.9939
| 1 | 5 | 280 | 256 | 246 | 271 | 275 | 267 | 277 | 264 | 274.9 | 258.2955 | 262 | 270.57 | 289.43 | 264.9 | 259 | 276.5535 | 266 | 264 | 282 | 246 | 270 | 241.74 | 291.4196 | 270 | 255.2 | 257 |  | 55
| 1 | 6 | 279 | 263 | 244 | 268 | 265 | 265 | 281 | 266 | 273.81 | 259.406 | 258 | 285.94 | 290.23 | 260.7 | 256 | 278.3287 | 268 | 255 | 263 | 245 | 264 | 239.94 | 294.4101 | 273 | 242.4 | 257
| Mean |  | 278.3333 | 262.3333 | 245.8333 | 272.5 | 265.1667 | 266.5 | 279.3333 | 267.1667 | 273.4883 | 262.7855 | 264.5 | 278.4433 | 289.7617 | 263.4 | 257.1667 | 281.3191 | 264.1667 | 262.8333 | 272.6667 | 242.3333 | 277.5 | 239.37 | 293.1365 | 272.3333 | 249.7833 | 258.8333
| Median |  | 279 | 259.5 | 245.5 | 271.5 | 265.5 | 266 | 278.5 | 267.5 | 273.24 | 261.4895 | 263.5 | 279.265 | 289.83 | 264 | 256.5 | 281.5554 | 265 | 264 | 274.5 | 242 | 275 | 240.84 | 294.1056 | 270 | 249.85 | 259
| Std Dev. |  | 2.8048 | 8.2381 | 2.3166 | 4.5497 | 6.1128 | 3.8341 | 4.9666 | 2.137 | 3.3648 | 4.4432 | 5.3572 | 6.0054 | 2.3017 | 1.9769 | 4.2622 | 3.7432 | 3.1885 | 4.446 | 7.7889 | 5.9217 | 13.6931 | 6.0319 | 2.7353 | 4.9261 | 4.5208 | 1.6021
| Rel.Std.Dev. |  | 1.008% | 3.14% | 0.9423% | 1.67% | 2.305% | 1.439% | 1.778% | 0.7999% | 1.23% | 1.691% | 2.025% | 2.157% | 0.7944% | 0.7505% | 1.657% | 1.331% | 1.207% | 1.692% | 2.857% | 2.444% | 4.934% | 2.52% | 0.9331% | 1.809% | 1.81% | 0.619%
| PDM3 |  | 4.639% | -1.376% | -7.579% | 2.446% | -0.311% | 0.1903% | 5.015% | 0.4409% | 2.818% | -1.206% | -0.5616% | 4.68% | 8.935% | -0.9752% | -3.319% | 5.761% | -0.687% | -1.188% | 2.509% | -8.895% | 4.326% | -10.01% | 10.2% | 2.383% | -6.094% | -2.692%
| Z-Score (Absolute) |  | 1.15 | 0.32 | 1.84 | 0.61 | 0.06 | 0.06 | 1.24 | 0.12 | 0.7 | 0.28 | 0.12 | 1.16 | 2.2 | 0.22 | 0.8 | 1.43 | 0.15 | 0.28 | 0.63 | 2.16 | 1.07 | 2.44 | 2.51 | 0.6 | 1.48 | 0.64
| Analytical results for Ta in OREAS 505b (Certified Value 1.05 ppm) | [28 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.29 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 1.02 | 1.4 | 0.9 | 0.97 | 0.98 | 1.06 | 1.03 | 1.01 | 1.11 | 1.06 | 1.01 | 1.1324 | 0.96 | 1.06 | 1.04 | < 5 | 1.1 | 1.496 | 1.1 | 1.16 | [7 empty] | 1
| 1 | 2 | 1.02 | 1.4 | 0.9 | 1.13 | 1.02 | 1.06 | 1.1 | 1.02 | 1.06 | 1.07 | 0.97 | 1.1038 | 1.02 | 1.08 | 1.02 | < 5 | 1 | 1.484 | 1.1 | 0.96 | [7 empty] | 28
| 1 | 3 | 1 | 1.3 | 0.9 | 1.12 | 0.99 | 1.06 | 1.1 | 1.06 | 1.07 | 1.16 | 0.97 | 1.0983 | 1.02 | 1.04 | 0.99 | < 5 | 1 | 1.444 | 1.1 | 1.06 | [7 empty] | 16
| 1 | 4 | 1.01 | 1.3 | 0.9 | 1.07 | 1.05 | 1.06 | 1.12 | 1.01 | 1.1 | 1.08 | 1.01 | 1.1103 | 0.96 | 1.04 | 1.03 | < 5 | 0.9 | 1.43 | 1 | 0.95 | [7 empty] | 1.0465
| 1 | 5 | 1.03 | 1.4 | 0.9 | 1.17 | 1.02 | 1.07 | 1.06 | 0.97 | 1.14 | 1.09 | 0.98 | 1.1077 | 0.99 | 1.07 | 1.04 | < 5 | 1 | 1.458 | 1 | 0.88 | [7 empty] | 56
| 1 | 6 | 1.1 | 1.4 | 0.9 | 1.04 | 0.97 | 1.02 | 1.07 | 1.04 | 1.09 | 1.04 | 1.07 | 1.134 | 1.03 | 1.09 | 1.04 | < 5 | 0.9 | 1.501 | 1.1 | 1.01
| Mean |  | 1.03 | 1.3667 | 0.9 | 1.0833 | 1.005 | 1.055 | 1.08 | 1.0183 | 1.095 | 1.0833 | 1.0017 | 1.1144 | 0.99667 | 1.0633 | 1.0267 |  | 0.98333 | 1.4688 | 1.0667 | 1.0033
| Median |  | 1.02 | 1.4 | 0.9 | 1.095 | 1.005 | 1.06 | 1.085 | 1.015 | 1.095 | 1.075 | 0.995 | 1.109 | 1.005 | 1.065 | 1.035 |  | 1 | 1.471 | 1.1 | 0.985
| Std Dev. |  | 0.035777 | 0.05164 | 0 | 0.072019 | 0.030166 | 0.017607 | 0.032863 | 0.030605 | 0.02881 | 0.041312 | 0.038166 | 0.015102 | 0.031411 | 0.020656 | 0.019664 |  | 0.075277 | 0.029137 | 0.05164 | 0.097707
| Rel.Std.Dev. |  | 3.474% | 3.779% | 0% | 6.648% | 3.002% | 1.669% | 3.043% | 3.005% | 2.631% | 3.813% | 3.81% | 1.355% | 3.152% | 1.943% | 1.915% |  | 7.655% | 1.984% | 4.841% | 9.738%
| PDM3 |  | -1.579% | 30.59% | -14% | 3.517% | -3.968% | 0.8095% | 3.198% | -2.694% | 4.632% | 3.517% | -4.287% | 6.485% | -4.764% | 1.606% | -1.898% |  | -6.039% | 40.35% | 1.924% | -4.127%
| Z-Score (Absolute) |  | 0.51 | NA | NA | 0.43 | 0.95 | 0.07 | 0.37 | 0.72 | 0.63 | 0.43 | 1.01 | 0.97 | 1.1 | 0.07 | 0.57 | 25.35 | NA | 7.21 | NA | 0.98
Results from laboratories 1.02, 1.03, 1.23 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tb in OREAS 505b (Certified Value 0.69 ppm) | [28 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab | [7 empty] | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.17; 1.06: 1.19; 1.07: 1.23; 1.08: 1.26; 1.09: 1.29 | [7 empty] | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS | [7 empty] | 2
2
| 1 | 1 | 1.01: 0.6; 1.02: 0.74; 1.03: 0.5305; 1.04: 0.78602; 1.05: 0.66; 1.06: 0.78; 1.07: 0.7; 1.08: 0.36546; 1.09: 0.69 | [7 empty] | 1
| 1 | 2 | 1.01: 0.6; 1.02: 0.74; 1.03: 0.6146; 1.04: 0.779; 1.05: 0.62; 1.06: 0.77; 1.07: 0.7; 1.08: 0.37132; 1.09: 0.64 | [7 empty] | 29
| 1 | 3 | 1.01: 0.7; 1.02: 0.76; 1.03: 0.5479; 1.04: 0.78806; 1.05: 0.6; 1.06: 0.74; 1.07: 0.7; 1.08: 0.35637; 1.09: 0.66 | [7 empty] | 16
| 1 | 4 | 1.01: 0.6; 1.02: 0.73; 1.03: 0.5549; 1.04: 0.79342; 1.05: 0.62; 1.06: 0.75; 1.07: 0.7; 1.08: 0.3018; 1.09: 0.68 | [7 empty] | 0.69477
| 1 | 5 | 1.01: 0.6; 1.02: 0.75; 1.03: 0.595; 1.04: 0.77859; 1.05: 0.66; 1.06: 0.75; 1.07: 0.7; 1.08: 0.29999; 1.09: 0.66 | [7 empty] | 57
| 1 | 6 | 1.01: 0.6; 1.02: 0.78; 1.03: 0.5787; 1.04: 0.76508; 1.05: 0.71; 1.06: 0.77; 1.07: 0.7; 1.08: 0.37904; 1.09: 0.64
| Mean |  | 1.01: 0.61667; 1.02: 0.75; 1.03: 0.57027; 1.04: 0.78169; 1.05: 0.645; 1.06: 0.76; 1.07: 0.7; 1.08: 0.34566; 1.09: 0.66167
| Median |  | 1.01: 0.6; 1.02: 0.745; 1.03: 0.5668; 1.04: 0.78251; 1.05: 0.64; 1.06: 0.76; 1.07: 0.7; 1.08: 0.36092; 1.09: 0.66
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.017889; 1.03: 0.031515; 1.04: 0.0098948; 1.05: 0.039875; 1.06: 0.015492; 1.07: 1.2162e-16; 1.08: 0.035466; 1.09: 0.020412
| Rel.Std.Dev. |  | 1.01: 6.62%; 1.02: 2.385%; 1.03: 5.526%; 1.04: 1.266%; 1.05: 6.182%; 1.06: 2.038%; 1.07: 1.737e-14%; 1.08: 10.26%; 1.09: 3.085%
| PDM3 |  | 1.01: -11.24%; 1.02: 7.949%; 1.03: -17.92%; 1.04: 12.51%; 1.05: -7.164%; 1.06: 9.389%; 1.07: 0.7526%; 1.08: -50.25%; 1.09: -4.765%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.65; 1.03: 0.67; 1.04: 0.89; 1.05: 0.12; 1.06: 0.73; 1.07: NA; 1.08: 2.33; 1.09: 0
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 505b (Certified Value 0.17 ppm) | [28 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 0.1 | < 0.5 | 0.18 | 0.18 | 0.17 | 0.2 | 0.18 | 0.15 | < 0.2 | < 5 | 0.89 | 0.16 | 0.15519 | 0.2 | 0.17 | 0.17 | 7.4 | 0.24 | < 10 | 0.243 | 0.1 | 0.18 | [5 empty] | 1
| 1 | 2 | 0.1 | < 0.5 | 0.23 | 0.12 | 0.16 | 0.19 | 0.18 | 0.19 | 0.2 | < 5 | 0.61 | 0.14 | 0.14945 | 0.21 | 0.17 | 0.18 | 6.9 | 0.24 | < 10 | 0.246 | 0.2 | 0.17 | [5 empty] | 30
| 1 | 3 | 0.2 | < 0.5 | 0.21 | 0.14 | 0.14 | 0.16 | 0.19 | 0.15 | 0.2 | < 5 | 0.94 | 0.13 | 0.21871 | 0.18 | 0.16 | 0.16 | 7.1 | 0.14 | < 10 | 0.245 | < 0.1 | 0.18 | [5 empty] | 16
| 1 | 4 | 0.1 | < 0.5 | 0.18 | 0.13 | 0.17 | 0.19 | 0.18 | 0.16 | < 0.2 | < 5 | 0.88 | 0.14 | 0.23213 | 0.2 | 0.16 | 0.21 | 7.3 | 0.14 | < 10 | 0.242 | 0.1 | 0.21 | [5 empty] | 0.17249
| 1 | 5 | 0.2 | < 0.5 | 0.21 | 0.12 | 0.17 | 0.18 | 0.18 | 0.15 | < 0.2 | < 5 | 0.37 | 0.14 | 0.17698 | 0.18 | 0.16 | 0.16 | 7.2 | 0.19 | < 10 | 0.24 | 0.2 | 0.16 | [5 empty] | 58
| 1 | 6 | 0.2 | < 0.5 | 0.2 | 0.14 | 0.2 | 0.16 | 0.21 | 0.15 | < 0.2 | < 5 | 0.51 | 0.16 | 0.13747 | 0.19 | 0.17 | 0.18 | 7.3 | 0.12 | < 10 | 0.245 | 0.1 | 0.2
| Mean |  | 0.15 |  | 0.20167 | 0.13833 | 0.16833 | 0.18 | 0.18667 | 0.15833 | 0.2 |  | 0.7 | 0.145 | 0.17832 | 0.19333 | 0.165 | 0.17667 | 7.2 | 0.17833 |  | 0.2435 | 0.14 | 0.18333
| Median |  | 0.15 |  | 0.205 | 0.135 | 0.17 | 0.185 | 0.18 | 0.15 | 0.2 |  | 0.745 | 0.14 | 0.16608 | 0.195 | 0.165 | 0.175 | 7.25 | 0.165 |  | 0.244 | 0.1 | 0.18
| Std Dev. |  | 0.054772 |  | 0.019408 | 0.022286 | 0.019408 | 0.016733 | 0.012111 | 0.016021 | 0 |  | 0.2363 | 0.012247 | 0.0389 | 0.012111 | 0.0054772 | 0.018619 | 0.17889 | 0.053072 |  | 0.0022583 | 0.054772 | 0.018619
| Rel.Std.Dev. |  | 36.51% |  | 9.624% | 16.11% | 11.53% | 9.296% | 6.488% | 10.12% | 0% |  | 33.76% | 8.447% | 21.81% | 6.264% | 3.32% | 10.54% | 2.485% | 29.76% |  | 0.9274% | 39.12% | 10.16%
| PDM3 |  | -13.04% |  | 16.92% | -19.8% | -2.408% | 4.356% | 8.221% | -8.205% | 15.95% |  | 305.8% | -15.94% | 3.383% | 12.09% | -4.34% | 2.424% | 4074% | 3.39% |  | 41.17% | -18.83% | 6.289%
| Z-Score (Absolute) |  | NA | 2.51 | 0.74 | 1.59 | 0.49 | 0.06 | 0.18 | 0.86 | NA | 85.27 | 19.06 | 1.35 | 0.12 | 0.43 | 0.61 | 0.18 | NA | 0.12 | 177.22 | 2.27 | NA | 0.06
Results from laboratories 1.01, 1.09, 1.22 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 505b (Certified Value 12.3 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 1
2
| 1 | 1 | 13.7 | 12.2 | 11.3 | 12.9 | 12.8 | 11.6 | 14.2 | 12.55 | 13.39 | 10.8158 | 12.2058 | 11.3 | 12.2 | 13.4257 | 12.19 | 13.25 | 12.7 | 12.1 | 11.38 | 11.9 | 11 | [6 empty] | 1
| 1 | 2 | 13.7 | 11.8 | 11.2 | 13.25 | 12.75 | 11.95 | 13.85 | 12.55 | 13 | 12.2085 | 12.2897 | 11.3 | 11.6 | 13.0431 | 11.9 | 12.45 | 13.05 | 12.9 | 10.899 | 11.8 | 9.7 | [6 empty] | 31
| 1 | 3 | 13.4 | 11.5 | 11.4 | 13.14 | 12.4 | 11.85 | 14.05 | 13.4 | 13.29 | 10.9255 | 13.9765 | 11.8 | 10.4 | 13.1988 | 12.57 | 12.05 | 12.15 | 12 | 11.15 | 12.8 | 10.4 | [6 empty] | 16
| 1 | 4 | 13.5 | 12.2 | 11.3 | 13.04 | 12.75 | 11.95 | 14.75 | 12.6 | 13.76 | 11.4179 | 13.4316 | 11.3 | 11.6 | 13.4201 | 12.14 | 12.55 | 13.35 | 12.5 | 11.227 | 12.4 | 10.7 | [6 empty] | 12.3089
| 1 | 5 | 13.6 | 11.4 | 11.2 | 12.9 | 12.95 | 12 | 13.6 | 12 | 13.74 | 11.6259 | 13.3851 | 11.6 | 11.4 | 13.1058 | 12.44 | 13.25 | 12.55 | 11.5 | 10.965 | 12.2 | 10.3 | [6 empty] | 59
| 1 | 6 | 13.5 | 11.8 | 11.4 | 13.64 | 13.45 | 11.7 | 14.15 | 12.65 | 12.51 | 11.3281 | 13.326 | 11.3 | 12.7 | 13.5084 | 12.29 | 11.7 | 12.65 | 11.9 | 11.367 | 12.7 | 9.7
| Mean |  | 13.5667 | 11.8167 | 11.3 | 13.145 | 12.85 | 11.8417 | 14.1 | 12.625 | 13.2817 | 11.387 | 13.1024 | 11.4333 | 11.65 | 13.2836 | 12.255 | 12.5417 | 12.7417 | 12.15 | 11.1647 | 12.3 | 10.3
| Median |  | 13.55 | 11.8 | 11.3 | 13.09 | 12.775 | 11.9 | 14.1 | 12.575 | 13.34 | 11.373 | 13.3555 | 11.3 | 11.6 | 13.3094 | 12.24 | 12.5 | 12.675 | 12.05 | 11.1885 | 12.3 | 10.35
| Std Dev. |  | 0.12111 | 0.33714 | 0.089443 | 0.27826 | 0.34496 | 0.15943 | 0.3873 | 0.44805 | 0.47444 | 0.5051 | 0.70237 | 0.21602 | 0.7791 | 0.19287 | 0.23586 | 0.62643 | 0.41523 | 0.48888 | 0.20092 | 0.40988 | 0.52536
| Rel.Std.Dev. |  | 0.8927% | 2.853% | 0.7915% | 2.117% | 2.685% | 1.346% | 2.747% | 3.549% | 3.572% | 4.436% | 5.361% | 1.889% | 6.688% | 1.452% | 1.925% | 4.995% | 3.259% | 4.024% | 1.8% | 3.332% | 5.101%
| PDM3 |  | 10.22% | -3.999% | -8.197% | 6.792% | 4.396% | -3.796% | 14.55% | 2.568% | 7.903% | -7.49% | 6.447% | -7.114% | -5.353% | 7.919% | -0.4382% | 1.891% | 3.516% | -1.291% | -9.296% | -0.07261% | -16.32%
| Z-Score (Absolute) |  | 1.06 | 0.41 | 0.84 | 0.71 | 0.46 | 0.39 | 1.51 | 0.27 | 0.82 | 0.77 | 0.67 | 0.73 | 0.55 | 0.83 | 0.04 | 0.2 | 0.37 | 0.13 | 0.95 | 0 | 1.68
| Analytical results for Ti in OREAS 505b (Certified Value 0.337 wt.%) | [28 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 3
3
| 1 | 1 | 0.343 | 0.332 | 0.34 | 0.34 | 0.324 | 0.332 | 0.336 | 0.334 | 0.3412 | 0.35243 | 0.35 | 0.32 | 0.33907 | 0.34 | 0.34 | 0.34803 | 0.35 | 0.326 | 0.328 | 0.32 | 0.338 | 0.32549 | 0.28927 | 0.347 | 0.31 | 0.321 |  | 1
| 1 | 2 | 0.34 | 0.323 | 0.349 | 0.34 | 0.339 | 0.337 | 0.343 | 0.341 | 0.3409 | 0.35273 | 0.35 | 0.3 | 0.33469 | 0.34 | 0.34 | 0.34257 | 0.34 | 0.329 | 0.338 | 0.32 | 0.336 | 0.31909 | 0.27948 | 0.341 | 0.31 | 0.323 |  | 32
| 1 | 3 | 0.339 | 0.331 | 0.348 | 0.34 | 0.337 | 0.338 | 0.332 | 0.337 | 0.3487 | 0.348 | 0.34 | 0.35 | 0.33095 | 0.34 | 0.35 | 0.34334 | 0.34 | 0.332 | 0.335 | 0.32 | 0.326 | 0.32205 | 0.29058 | 0.358 | 0.31 | 0.324 |  | 16
| 1 | 4 | 0.337 | 0.334 | 0.342 | 0.34 | 0.336 | 0.339 | 0.352 | 0.345 | 0.3468 | 0.34322 | 0.35 | 0.33 | 0.32893 | 0.34 | 0.34 | 0.34186 | 0.35 | 0.331 | 0.341 | 0.32 | 0.337 | 0.3225 | 0.28744 | 0.345 | 0.31 | 0.324 |  | 0.33719
| 1 | 5 | 0.333 | 0.324 | 0.346 | 0.34 | 0.343 | 0.337 | 0.335 | 0.334 | 0.35 | 0.34406 | 0.35 | 0.33 | 0.33045 | 0.33 | 0.34 | 0.34942 | 0.34 | 0.333 | 0.349 | 0.32 | 0.347 | 0.32513 | 0.28341 | 0.343 | 0.32 | 0.318 |  | 60
| 1 | 6 | 0.332 | 0.34 | 0.343 | 0.34 | 0.333 | 0.33 | 0.348 | 0.334 | 0.3517 | 0.35462 | 0.35 | 0.33 | 0.33425 | 0.34 | 0.34 | 0.34595 | 0.34 | 0.322 | 0.327 | 0.32 | 0.345 | 0.32538 | 0.28544 | 0.346 | 0.3 | 0.321
| Mean |  | 0.33733 | 0.33067 | 0.34467 | 0.34 | 0.33533 | 0.3355 | 0.341 | 0.3375 | 0.34655 | 0.34918 | 0.34833 | 0.32667 | 0.33306 | 0.33833 | 0.34167 | 0.34519 | 0.34333 | 0.32883 | 0.33633 | 0.32 | 0.33817 | 0.32327 | 0.28594 | 0.34667 | 0.31 | 0.32183
| Median |  | 0.338 | 0.3315 | 0.3445 | 0.34 | 0.3365 | 0.337 | 0.3395 | 0.3355 | 0.34775 | 0.35021 | 0.35 | 0.33 | 0.3326 | 0.34 | 0.34 | 0.34465 | 0.34 | 0.33 | 0.3365 | 0.32 | 0.3375 | 0.32381 | 0.28644 | 0.3455 | 0.31 | 0.322
| Std Dev. |  | 0.0042269 | 0.006377 | 0.003559 | 0 | 0.0064704 | 0.0036194 | 0.0079498 | 0.0045935 | 0.0045531 | 0.004815 | 0.0040825 | 0.01633 | 0.0037018 | 0.0040825 | 0.0040825 | 0.003095 | 0.005164 | 0.0041673 | 0.0082865 | 0 | 0.0074677 | 0.0025467 | 0.0040804 | 0.0059554 | 0.0063246 | 0.0023166
| Rel.Std.Dev. |  | 1.253% | 1.929% | 1.033% | 0% | 1.93% | 1.079% | 2.331% | 1.361% | 1.314% | 1.379% | 1.172% | 4.999% | 1.111% | 1.207% | 1.195% | 0.8966% | 1.504% | 1.267% | 2.464% | 0% | 2.208% | 0.7878% | 1.427% | 1.718% | 2.04% | 0.7198%
| PDM3 |  | 0.04349% | -1.934% | 2.218% | 0.8343% | -0.5497% | -0.5002% | 1.131% | 0.09292% | 2.777% | 3.556% | 3.306% | -3.12% | -1.225% | 0.3401% | 1.329% | 2.375% | 1.823% | -2.477% | -0.2531% | -5.097% | 0.2906% | -4.127% | -15.2% | 2.811% | -8.063% | -4.553%
| Z-Score (Absolute) |  | 0.01 | 0.65 | 0.7 | 0.25 | 0.2 | 0.18 | 0.35 | 0.01 | 0.88 | 1.13 | 1.05 | 1.04 | 0.42 | 0.09 | 0.41 | 0.75 | 0.57 | 0.83 | 0.1 | 1.68 | 0.07 | 1.36 | 4.96 | 0.89 | 2.64 | 1.5
| Analytical results for Tl in OREAS 505b (Certified Value 0.82 ppm) | [28 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 0.81 | 0.8 | 0.81 | 0.83 | 0.78 | 0.79 | 0.83 | 0.79 | 0.88 | 0.83426 | 0.6 | < 0.1 | 0.83 | 0.85581 | 0.59 | 0.83 | 0.76 | 0.76 | 0.875 | 0.8 | 0.77 | [6 empty] | 1
| 1 | 2 | 0.85 | 0.8 | 0.78 | 0.84 | 0.74 | 0.78 | 0.85 | 0.78 | 0.87 | 0.84979 | 0.58 | < 0.1 | 0.82 | 0.81815 | 0.59 | 0.84 | 0.76 | 0.75 | 0.864 | 0.9 | 0.75 | [6 empty] | 33
| 1 | 3 | 0.84 | 0.8 | 0.83 | 0.87 | 0.77 | 0.77 | 0.86 | 0.83 | 0.85 | 0.90122 | 0.6 | 0.17 | 0.8 | 0.85259 | 0.54 | 0.86 | 0.74 | 0.77 | 0.895 | 0.8 | 0.68 | [6 empty] | 16
| 1 | 4 | 0.81 | 0.8 | 0.78 | 0.81 | 0.81 | 0.81 | 0.88 | 0.79 | 0.91 | 0.90563 | 0.6 | 0.76 | 0.81 | 0.83626 | 0.54 | 0.82 | 0.79 | 0.75 | 0.899 | 0.7 | 0.76 | [6 empty] | 0.81893
| 1 | 5 | 0.8 | 0.8 | 0.81 | 0.84 | 0.83 | 0.81 | 0.82 | 0.79 | 0.86 | 0.91147 | 0.59 | < 0.1 | 0.78 | 0.80897 | 0.52 | 0.84 | 0.82 | 0.77 | 0.868 | 0.8 | 0.76 | [6 empty] | 61
| 1 | 6 | 0.82 | 0.8 | 0.8 | 0.87 | 0.77 | 0.77 | 0.84 | 0.85 | 0.87 | 0.90645 | 0.58 | < 0.1 | 0.85 | 0.862 | 0.57 | 0.8 | 0.74 | 0.78 | 0.893 | 0.9 | 0.73
| Mean |  | 0.82167 | 0.8 | 0.80167 | 0.84333 | 0.78333 | 0.78833 | 0.84667 | 0.805 | 0.87333 | 0.88481 | 0.59167 | 0.465 | 0.815 | 0.83896 | 0.55833 | 0.83167 | 0.76833 | 0.76333 | 0.88233 | 0.81667 | 0.74167
| Median |  | 0.815 | 0.8 | 0.805 | 0.84 | 0.775 | 0.785 | 0.845 | 0.79 | 0.87 | 0.90342 | 0.595 | 0.465 | 0.815 | 0.84443 | 0.555 | 0.835 | 0.76 | 0.765 | 0.884 | 0.8 | 0.755
| Std Dev. |  | 0.019408 | 1.2162e-16 | 0.019408 | 0.023381 | 0.032042 | 0.018348 | 0.021602 | 0.028107 | 0.020656 | 0.033654 | 0.0098319 | 0.41719 | 0.02429 | 0.021636 | 0.029269 | 0.020412 | 0.031252 | 0.012111 | 0.015148 | 0.075277 | 0.033116
| Rel.Std.Dev. |  | 2.362% | 1.52e-14% | 2.421% | 2.772% | 4.09% | 2.328% | 2.551% | 3.492% | 2.365% | 3.804% | 1.662% | 89.72% | 2.98% | 2.579% | 5.242% | 2.454% | 4.067% | 1.587% | 1.717% | 9.218% | 4.465%
| PDM3 |  | 0.3342% | -2.312% | -2.108% | 2.98% | -4.347% | -3.736% | 3.387% | -1.701% | 6.643% | 8.044% | -27.75% | -43.22% | -0.4799% | 2.446% | -31.82% | 1.555% | -6.178% | -6.789% | 7.742% | -0.2764% | -9.435%
| Z-Score (Absolute) |  | 0.29 | NA | 0.06 | 0.67 | 0.38 | 0.29 | 0.73 | 0 | 1.2 | 1.4 | 3.75 | 10.85 | 0.18 | 0.6 | 4.34 | 0.47 | 0.64 | 0.73 | 1.36 | NA | 1.11
Results from laboratories 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 505b (Certified Value 0.18 ppm) | [28 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.19; 1.06: 1.23; 1.07: 1.26 | [6 empty] | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.139; 1.04: 0.20577; 1.05: 0.2; 1.06: 0.2; 1.07: 0.174 | [6 empty] | 1
| 1 | 2 | 1.01: 0.2; 1.02: 0.19; 1.03: 0.156; 1.04: 0.20154; 1.05: 0.21; 1.06: 0.2; 1.07: 0.17 | [6 empty] | 8
| 1 | 3 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.1434; 1.04: 0.19671; 1.05: 0.18; 1.06: 0.2; 1.07: 0.175 | [6 empty] | 16
| 1 | 4 | 1.01: 0.2; 1.02: 0.17; 1.03: 0.1378; 1.04: 0.20188; 1.05: 0.2; 1.06: 0.2; 1.07: 0.171 | [6 empty] | 0.17929
| 1 | 5 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.1448; 1.04: 0.20647; 1.05: 0.19; 1.06: 0.2; 1.07: 0.167 | [6 empty] | 62
| 1 | 6 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.1489; 1.04: 0.20356; 1.05: 0.19; 1.06: 0.2; 1.07: 0.186
| Mean |  | 1.01: 0.2; 1.02: 0.18; 1.03: 0.14498; 1.04: 0.20265; 1.05: 0.195; 1.06: 0.2; 1.07: 0.17383
| Median |  | 1.01: 0.2; 1.02: 0.18; 1.03: 0.1441; 1.04: 0.20272; 1.05: 0.195; 1.06: 0.2; 1.07: 0.1725
| Std Dev. |  | 1.01: 3.0405e-17; 1.02: 0.0063246; 1.03: 0.0067327; 1.04: 0.0035275; 1.05: 0.010488; 1.06: 3.0405e-17; 1.07: 0.0066156
| Rel.Std.Dev. |  | 1.01: 1.52e-14%; 1.02: 3.514%; 1.03: 4.644%; 1.04: 1.741%; 1.05: 5.379%; 1.06: 1.52e-14%; 1.07: 3.806%
| PDM3 |  | 1.01: 11.55%; 1.02: 0.3937%; 1.03: -19.14%; 1.04: 13.03%; 1.05: 8.76%; 1.06: 11.55%; 1.07: -3.046%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0; 1.03: 1.57; 1.04: 1.02; 1.05: 0.67; 1.06: NA; 1.07: 0.28
Results from laboratories 1.03 and 1.23 were removed due to their 0.1 ppm reading resolution.
| Analytical results for U in OREAS 505b (Certified Value 3.4 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 3.7 | 2.9 | 3 | 3.18 | 3.4 | 3.3 | 3.4 | 3.5 | 3.5 | 2.8978 | 3.2354 | 2.8 | 3.35 | 3.563 | 3.4 | 3.6 | 2.9 | 3.6 | 3.7475 | 3.7 | 3.5 | [6 empty] | 1
| 1 | 2 | 3.51 | 3.5 | 2.9 | 3.42 | 4 | 3.3 | 3.8 | 3.5 | 3.53 | 2.961 | 3.6624 | 2.84 | 3.07 | 3.9287 | 3.5 | 3.5 | 3.2 | 3.6 | 3.8905 | 3.4 | 3.3 | [6 empty] | 35
| 1 | 3 | 3.8 | 3.1 | 3.5 | 3.8 | 3.5 | 3.1 | 3.5 | 3.5 | 3.66 | 3.0425 | 3.3047 | 3.29 | 3.04 | 3.9274 | 3.1 | 3.6 | 2.8 | 3.8 | 4.6622 | 3.9 | 3.4 | [6 empty] | 16
| 1 | 4 | 3.84 | 3.6 | 3.1 | 3.93 | 3.3 | 3.2 | 4.6 | 3.6 | 3.56 | 2.9016 | 3.4751 | 2.87 | 4.09 | 3.8428 | 3.2 | 3.4 | 3.1 | 4 | 2.88 | 4.1 | 3.3 | [6 empty] | 3.3984
| 1 | 5 | 3.8 | 3.2 | 3 | 3.86 | 3.1 | 3.7 | 3.6 | 3.2 | 3.89 | 2.7701 | 3.0972 | 3.29 | 3.15 | 3.6487 | 3.3 | 4 | 3.3 | 3.1 | 3.8792 | 3.3 | 3.1 | [6 empty] | 63
| 1 | 6 | 3.6 | 3.2 | 2.9 | 3.97 | 3.6 | 2.9 | 3.4 | 3.8 | 3.56 | 2.7456 | 3.4897 | 2.8 | 3.69 | 3.7003 | 3.2 | 3.3 | 3.3 | 3.4 | 2.8753 | 3.6 | 3.3
| Mean |  | 3.7083 | 3.25 | 3.0667 | 3.6933 | 3.4833 | 3.25 | 3.7167 | 3.5167 | 3.6167 | 2.8864 | 3.3774 | 2.9817 | 3.3983 | 3.7685 | 3.2833 | 3.5667 | 3.1 | 3.5833 | 3.6558 | 3.6667 | 3.3167
| Median |  | 3.75 | 3.2 | 3 | 3.83 | 3.45 | 3.25 | 3.55 | 3.5 | 3.56 | 2.8997 | 3.3899 | 2.855 | 3.25 | 3.7716 | 3.25 | 3.55 | 3.15 | 3.6 | 3.8133 | 3.65 | 3.3
| Std Dev. |  | 0.1306 | 0.25884 | 0.22509 | 0.31935 | 0.30605 | 0.26646 | 0.45789 | 0.19408 | 0.14431 | 0.11282 | 0.20361 | 0.24028 | 0.41562 | 0.15341 | 0.1472 | 0.24221 | 0.20976 | 0.31252 | 0.68372 | 0.30111 | 0.13292
| Rel.Std.Dev. |  | 3.522% | 7.964% | 7.34% | 8.647% | 8.786% | 8.199% | 12.32% | 5.519% | 3.99% | 3.909% | 6.029% | 8.059% | 12.23% | 4.071% | 4.483% | 6.791% | 6.767% | 8.721% | 18.7% | 8.212% | 4.008%
| PDM3 |  | 9.119% | -4.368% | -9.762% | 8.677% | 2.498% | -4.368% | 9.364% | 3.479% | 6.421% | -15.07% | -0.6185% | -12.26% | -0.003029% | 10.89% | -3.387% | 4.95% | -8.782% | 5.441% | 7.573% | 7.893% | -2.406%
| Z-Score (Absolute) |  | 0.76 | 0.79 | 1.4 | 0.71 | 0 | 0.79 | 0.79 | 0.11 | 0.45 | 2.01 | 0.36 | 1.69 | 0.29 | 0.96 | 0.67 | 0.28 | 1.29 | 0.34 | 0.58 | 0.62 | 0.56
| Analytical results for V in OREAS 505b (Certified Value 82 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 78 | 85 | 82 | 83 | 76 | 82 | 80 | 81 | 83 | 80.8932 | 83.6745 | 71.4 | 82.49 | 81.8815 | 84.31 | 78 | 79.7845 | 87 | 79 | 83 | 82.6 | 80 | 84.89 | 72.7649 | 85 | 80 | 81 | 1
| 1 | 2 | 78 | 82 | 82 | 83 | 81 | 81 | 82 | 82 | 84 | 81.5245 | 84.4585 | 73.5 | 76.15 | 81.3174 | 84.57 | 76 | 79.634 | 85 | 80 | 86 | 81.6 | 81 | 85.89 | 71.0661 | 84 | 82 | 82 | 17
| 1 | 3 | 77 | 84 | 82 | 82 | 80 | 82 | 81 | 83 | 85 | 80.4173 | 81.9553 | 72.5 | 85.75 | 81.5084 | 82.9 | 74 | 80.8408 | 76 | 81 | 85 | 81.4 | 82 | 84.74 | 72.7642 | 87 | 81 | 81 | 16
| 1 | 4 | 78 | 85 | 81 | 82 | 80 | 81 | 85 | 83 | 83 | 80.4043 | 83.3231 | 75.4 | 83.27 | 81.2159 | 83.54 | 78 | 80.9241 | 82 | 82 | 87 | 82.7 | 80 | 85.5 | 73.4742 | 84 | 82 | 81 | 82.0769
| 1 | 5 | 78 | 82 | 83 | 83 | 82 | 80 | 81 | 80 | 79 | 80.7167 | 83.7715 | 71.4 | 84.34 | 85.323 | 83.1 | 78 | 80.9425 | 81 | 80 | 88 | 80.8 | 80 | 87.01 | 71.9658 | 84 | 82 | 80 | 64
| 1 | 6 | 75 | 87 | 83 | 81 | 79 | 80 | 82 | 82 | 87 | 80.3714 | 83.1555 | 72.9 | 83.3 | 81.6518 | 82.5 | 76 | 80.8568 | 76 | 80 | 83 | 83.7 | 81 | 86.81 | 72.7638 | 85 | 79 | 80
| Mean |  | 77.3333 | 84.1667 | 82.1667 | 82.3333 | 79.6667 | 81 | 81.8333 | 81.8333 | 83.5 | 80.7212 | 83.3897 | 72.85 | 82.55 | 82.1497 | 83.4867 | 76.6667 | 80.4971 | 81.1667 | 80.3333 | 85.3333 | 82.1333 | 80.6667 | 85.8067 | 72.4665 | 84.8333 | 81 | 80.8333
| Median |  | 78 | 84.5 | 82 | 82.5 | 80 | 81 | 81.5 | 82 | 83.5 | 80.567 | 83.4988 | 72.7 | 83.285 | 81.5801 | 83.32 | 77 | 80.8488 | 81.5 | 80 | 85.5 | 82.1 | 80.5 | 85.695 | 72.764 | 84.5 | 81.5 | 81
| Std Dev. |  | 1.2111 | 1.9408 | 0.75277 | 0.8165 | 2.0656 | 0.89443 | 1.7224 | 1.169 | 2.6646 | 0.44498 | 0.83501 | 1.5003 | 3.3313 | 1.5726 | 0.81505 | 1.633 | 0.61334 | 4.535 | 1.0328 | 2.0656 | 1.0577 | 0.8165 | 0.95223 | 0.83586 | 1.169 | 1.2649 | 0.75277
| Rel.Std.Dev. |  | 1.566% | 2.306% | 0.9162% | 0.9917% | 2.593% | 1.104% | 2.105% | 1.429% | 3.191% | 0.5512% | 1.001% | 2.059% | 4.036% | 1.914% | 0.9763% | 2.13% | 0.7619% | 5.587% | 1.286% | 2.421% | 1.288% | 1.012% | 1.11% | 1.153% | 1.378% | 1.562% | 0.9313%
| PDM3 |  | -5.779% | 2.546% | 0.1093% | 0.3124% | -2.937% | -1.312% | -0.2968% | -0.2968% | 1.734% | -1.652% | 1.599% | -11.24% | 0.5764% | 0.08862% | 1.718% | -6.592% | -1.925% | -1.109% | -2.124% | 3.967% | 0.06871% | -1.718% | 4.544% | -11.71% | 3.358% | -1.312% | -1.515%
| Z-Score (Absolute) |  | 2.27 | 1.18 | 0.17 | 0.25 | 1.09 | 0.42 | 0 | 0 | 0.84 | 0.56 | 0.79 | 4.53 | 0.36 | 0.16 | 0.83 | 2.61 | 0.67 | 0.34 | 0.76 | 1.77 | 0.15 | 0.59 | 2.01 | 4.73 | 1.51 | 0.42 | 0.5
| Analytical results for W in OREAS 505b (Certified Value 9.12 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.29 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 2
3
| 1 | 1 | 8 | 8.5 | 9.1 | 8.9 | 8.1 | 8.8 | 9.4 | 9.3 | 9.7 | 7.2663 | 9.5938 | 9.3 | < 5 | 9 | 9.9756 | < 20 | 9.6 | 8.1 | 10.7 | 8.3 | 10.224 | 7.1 | 8.7 | < 10 |  |  |  | 1
| 1 | 2 | 8.5 | 8.4 | 8.6 | 9.7 | 8.6 | 8.6 | 9.7 | 8.6 | 8.3 | 7.6176 | 9.6424 | 12.9 | < 5 | 9 | 9.6746 | < 20 | 10.9 | 7.7 | 10.2 | 9.5 | 10.64 | 8.8 | 8.5 | < 10 |  |  |  | 37
| 1 | 3 | 8.1 | 9.3 | 8.8 | 8.9 | 8.4 | 8 | 10.3 | 8.1 | 9.1 | 6.6142 | 9.7895 | 9.5 | < 5 | 9 | 10.3424 | < 20 | 8.4 | 7.8 | 10.2 | 7.8 | 10.638 | 9 | 7.8 | < 10 |  |  |  | 16
| 1 | 4 | 8.5 | 8.3 | 7.7 | 9 | 8.7 | 8.4 | 10 | 9.6 | 9.4 | 7.8337 | 10.2255 | 9.7 | < 5 | 10 | 9.1545 | < 20 | 8.1 | 8.1 | 10.1 | 7.9 | 10.429 | 8.8 | 8.6 | < 10 |  |  |  | 9.1165
| 1 | 5 | 8.6 | 8.4 | 8.2 | 8.7 | 10.6 | 9.7 | 9 | 9.3 | 8.9 | 6.7191 | 9.6161 | 9.1 | < 5 | 9 | 9.3531 | < 20 | 9.3 | 8.7 | 10.1 | 8.4 | 10.51 | 9.1 | 8.6 | 10 |  |  |  | 65
| 1 | 6 | 8.4 | 8.8 | 7.9 | 9 | 8.8 | 7.9 | 9.7 | 9.1 | 9.5 | 6.2483 | 10.2975 | 9.4 | < 5 | 10 | 10.2809 | < 20 | 10.2 | 8.4 | 10.4 | 8.8 | 10.411 | 9 | 8.3 | < 10
| Mean |  | 8.35 | 8.6167 | 8.3833 | 9.0333 | 8.8667 | 8.5667 | 9.6833 | 9 | 9.15 | 7.0499 | 9.8608 | 9.9833 |  | 9.3333 | 9.7968 |  | 9.4167 | 8.1333 | 10.2833 | 8.45 | 10.4753 | 8.6333 | 8.4167 | 10
| Median |  | 8.45 | 8.45 | 8.4 | 8.95 | 8.65 | 8.5 | 9.7 | 9.2 | 9.25 | 6.9927 | 9.716 | 9.45 |  | 9 | 9.8251 |  | 9.45 | 8.1 | 10.2 | 8.35 | 10.4695 | 8.9 | 8.55 | 10
| Std Dev. |  | 0.2429 | 0.37639 | 0.54191 | 0.34448 | 0.88468 | 0.6532 | 0.4535 | 0.55136 | 0.50498 | 0.62053 | 0.31863 | 1.4428 |  | 0.5164 | 0.48751 |  | 1.061 | 0.37238 | 0.23166 | 0.62849 | 0.15763 | 0.7607 | 0.33116
| Rel.Std.Dev. |  | 2.909% | 4.368% | 6.464% | 3.813% | 9.978% | 7.625% | 4.683% | 6.126% | 5.519% | 8.802% | 3.231% | 14.45% |  | 5.533% | 4.976% |  | 11.27% | 4.578% | 2.253% | 7.438% | 1.505% | 8.811% | 3.935%
| PDM3 |  | -8.408% | -5.483% | -8.042% | -0.9121% | -2.74% | -6.031% | 6.218% | -1.278% | 0.3677% | -22.67% | 8.164% | 9.509% |  | 2.379% | 7.463% |  | 3.293% | -10.78% | 12.8% | -7.311% | 14.91% | -5.3% | -7.676% | 9.691%
| Z-Score (Absolute) |  | 0.63 | 0.31 | 0.59 | 0.2 | 0 | 0.37 | 1 | 0.16 | 0.35 | 2.23 | 1.22 | 1.37 | 7.81 | NA | 1.14 | 1.39 | 0.67 | 0.9 | 1.74 | 0.51 | 1.97 | 0.29 | 0.55 | 3.72
Results from laboratory 1.17 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 505b (Certified Value 14.9 ppm) | [28 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  |  |  | 1
2
| 1 | 1 | 15.9 | 15.5 | 13.5 | 14.5 | 15.1 | 15.3 | 15.9 | 14.8 | 15.63 | 11.9962 | 15.739 | 17 | 15.1 | 17.0337 | 14.6 | 15.1 | 14.4 | 15 | 14.7 | 13.412 | 11.981 | 14.7 | 14.4 | 13.9 |  |  |  | 1
| 1 | 2 | 16.2 | 14.9 | 13.8 | 14.4 | 15 | 16 | 16 | 15 | 15.07 | 13.0351 | 15.1692 | 16.4 | 13.8 | 16.4749 | 14.7 | 14.1 | 14.5 | 12.9 | 15.4 | 13.788 | 11.7014 | 13.4 | 13 | 13.8 |  |  |  | 38
| 1 | 3 | 16 | 14.8 | 13.7 | 14.3 | 15.6 | 15.6 | 16.7 | 15.2 | 15.41 | 12.549 | 15.3556 | 17 | 14.2 | 16.7988 | 13.6 | 14.2 | 13.9 | 14.6 | 14.7 | 12.857 | 11.7713 | 16.1 | 13.8 | 13.6 |  |  |  | 16
| 1 | 4 | 15.8 | 15.8 | 14 | 14.9 | 15.3 | 15.8 | 16.6 | 15.1 | 15.21 | 11.9549 | 15.3461 | 16.5 | 14.7 | 16.9602 | 14.3 | 14.5 | 14.6 | 14.5 | 14.9 | 13.473 | 11.9311 | 14.8 | 14.7 | 14.2 |  |  |  | 14.8557
| 1 | 5 | 15.7 | 15.7 | 13.8 | 14 | 15.9 | 16 | 16 | 14.8 | 16.2 | 12.9511 | 15.3245 | 16.5 | 13.6 | 16.4501 | 14.5 | 14.7 | 14.6 | 14.5 | 14.7 | 13.552 | 11.6514 | 15.4 | 13.9 | 14 |  |  |  | 66
| 1 | 6 | 15.7 | 14.6 | 13.9 | 13.8 | 15.5 | 15.2 | 16.2 | 14.9 | 15.44 | 12.9515 | 15.3549 | 17.2 | 14.8 | 16.9124 | 13.8 | 14.5 | 14.2 | 15.1 | 14.6 | 13.62 | 11.8012 | 15.8 | 13.2 | 13.6
| Mean |  | 15.8833 | 15.2167 | 13.7833 | 14.3167 | 15.4 | 15.65 | 16.2333 | 14.9667 | 15.4933 | 12.573 | 15.3815 | 16.7667 | 14.3667 | 16.7717 | 14.25 | 14.5167 | 14.3667 | 14.4333 | 14.8333 | 13.4503 | 11.8062 | 15.0333 | 13.8333 | 13.85
| Median |  | 15.85 | 15.2 | 13.8 | 14.35 | 15.4 | 15.7 | 16.1 | 14.95 | 15.425 | 12.75 | 15.3505 | 16.75 | 14.45 | 16.8556 | 14.4 | 14.5 | 14.45 | 14.55 | 14.7 | 13.5125 | 11.7863 | 15.1 | 13.85 | 13.85
| Std Dev. |  | 0.19408 | 0.51153 | 0.17224 | 0.38687 | 0.33466 | 0.34496 | 0.33862 | 0.1633 | 0.39682 | 0.49297 | 0.18909 | 0.33862 | 0.59554 | 0.25147 | 0.45056 | 0.36009 | 0.27325 | 0.79415 | 0.29439 | 0.31849 | 0.12832 | 0.96885 | 0.65929 | 0.23452
| Rel.Std.Dev. |  | 1.222% | 3.362% | 1.25% | 2.702% | 2.173% | 2.204% | 2.086% | 1.091% | 2.561% | 3.921% | 1.229% | 2.02% | 4.145% | 1.499% | 3.162% | 2.481% | 1.902% | 5.502% | 1.985% | 2.368% | 1.087% | 6.445% | 4.766% | 1.693%
| PDM3 |  | 6.917% | 2.43% | -7.219% | -3.629% | 3.664% | 5.347% | 9.273% | 0.7467% | 4.292% | -15.37% | 3.539% | 12.86% | -3.292% | 12.9% | -4.077% | -2.282% | -3.292% | -2.843% | -0.1508% | -9.46% | -20.53% | 1.196% | -6.882% | -6.77%
| Z-Score (Absolute) |  | 1.06 | 0.47 | 0.78 | 0.31 | 0.63 | 0.85 | 1.36 | 0.25 | 0.72 | 1.84 | 0.62 | 1.83 | 0.27 | 1.83 | 0.37 | 0.14 | 0.27 | 0.21 | 0.14 | 1.07 | 2.51 | 0.31 | 0.74 | 0.72
| Analytical results for Yb in OREAS 505b (Certified Value 1.17 ppm) | [28 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  |  | 1
| No | No | 1.01: 1.03; 1.02: 1.04; 1.03: 1.10; 1.04: 1.11; 1.05: 1.17; 1.06: 1.19; 1.07: 1.23; 1.08: 1.26; 1.09: 1.29 |  |  |  | ppm
|  |  | 1.01: 4A*OES/MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*OES/MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 1.01: 1.1; 1.02: 1.2; 1.03: 0.9189; 1.04: 1.3835; 1.05: 1.3; 1.06: 1.23; 1.07: 1.1; 1.08: 1.0498; 1.09: 1.3 |  |  |  | 1
| 1 | 2 | 1.01: 1.1; 1.02: 1.1; 1.03: 1.0172; 1.04: 1.3788; 1.05: 1.3; 1.06: 1.23; 1.07: 1.3; 1.08: 1.0303; 1.09: 1.1 |  |  |  | 39
| 1 | 3 | 1.01: 1.1; 1.02: 1.2; 1.03: 0.9434; 1.04: 1.4031; 1.05: 1.3; 1.06: 1.19; 1.07: 1.2; 1.08: 1.03; 1.09: 1.1 |  |  |  | 16
| 1 | 4 | 1.01: 1; 1.02: 1.2; 1.03: 0.9207; 1.04: 1.3199; 1.05: 1.2; 1.06: 1.24; 1.07: 1.2; 1.08: 1.1002; 1.09: 1.3 |  |  |  | 1.1706
| 1 | 5 | 1.01: 1; 1.02: 1.2; 1.03: 0.9613; 1.04: 1.3734; 1.05: 1.2; 1.06: 1.2; 1.07: 1.2; 1.08: 1.0703; 1.09: 1.2 |  |  |  | 67
| 1 | 6 | 1.01: 1.1; 1.02: 1.2; 1.03: 1.0043; 1.04: 1.3923; 1.05: 1.3; 1.06: 1.18; 1.07: 1.2; 1.08: 1.0802; 1.09: 1.2
| Mean |  | 1.01: 1.0667; 1.02: 1.1833; 1.03: 0.96097; 1.04: 1.3752; 1.05: 1.2667; 1.06: 1.2117; 1.07: 1.2; 1.08: 1.0601; 1.09: 1.2
| Median |  | 1.01: 1.1; 1.02: 1.2; 1.03: 0.95235; 1.04: 1.3811; 1.05: 1.3; 1.06: 1.215; 1.07: 1.2; 1.08: 1.06; 1.09: 1.2
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.040825; 1.03: 0.041808; 1.04: 0.02904; 1.05: 0.05164; 1.06: 0.024833; 1.07: 0.063246; 1.08: 0.028325; 1.09: 0.089443
| Rel.Std.Dev. |  | 1.01: 4.841%; 1.02: 3.45%; 1.03: 4.351%; 1.04: 2.112%; 1.05: 4.077%; 1.06: 2.049%; 1.07: 5.27%; 1.08: 2.672%; 1.09: 7.454%
| PDM3 |  | 1.01: -8.881%; 1.02: 1.085%; 1.03: -17.91%; 1.04: 17.47%; 1.05: 8.204%; 1.06: 3.506%; 1.07: 2.509%; 1.08: -9.44%; 1.09: 2.509%
| Z-Score (Absolute) |  | 1.01: 1.35; 1.02: 0.17; 1.03: 2.42; 1.04: 1.77; 1.05: 0.67; 1.06: 0.12; 1.07: 0; 1.08: 1.41; 1.09: 0
| Analytical results for Zn in OREAS 505b (Certified Value 98 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 0
1
| 1 | 1 | 96 | 97 | 94.2 | 99 | 94 | 100 | 99 | 99 | 101 | 90.3521 | 101.8998 | 98.5 | 71.13 | 85.52 | 102.9 | 96 | 100.6294 | 99 | 97 | 105 | 98 | 107.1 | 93.748 | 92.772 | 84 | 104 | 91 | 1
| 1 | 2 | 100 | 94 | 95.4 | 99 | 98 | 103 | 101 | 101 | 101 | 90.473 | 101.4612 | 102 | 63.84 | 86.9216 | 107.3 | 96 | 100.2849 | 99 | 99 | 109 | 100 | 108.1 | 92.265 | 89.3197 | 81 | 95 | 91 | 18
| 1 | 3 | 95 | 97 | 94.8 | 98 | 97 | 101 | 99 | 100 | 102 | 90.4456 | 98.6811 | 97.3 | 72.99 | 94.8416 | 105.7 | 95 | 100.9418 | 97 | 99 | 108 | 100 | 107.1 | 92.299 | 92.1105 | 90 | 98 | 90 | 16
| 1 | 4 | 98 | 98 | 93.9 | 97 | 98 | 102 | 104 | 101 | 105 | 89.5202 | 98.7297 | 99.8 | 67.93 | 92.7392 | 103 | 94 | 101.4539 | 98 | 99 | 110 | 101 | 105 | 92.552 | 91.7233 | 90 | 104 | 93 | 98.2669
| 1 | 5 | 98 | 95 | 94 | 98 | 100 | 102 | 99 | 97 | 102 | 91.2781 | 99.0841 | 99 | 67.64 | 93.9968 | 105.8 | 95 | 98.9536 | 99 | 99 | 112 | 98 | 108.1 | 94.126 | 91.0651 | 88 | 101 | 88 | 68
| 1 | 6 | 97 | 98 | 96.2 | 98 | 97 | 100 | 102 | 98 | 103 | 91.2576 | 99.3486 | 98.6 | 68.02 | 102.2624 | 101.1 | 96 | 99.3383 | 99 | 96 | 106 | 101 | 107.1 | 94.127 | 91.9085 | 86 | 94 | 90
| Mean |  | 97.3333 | 96.5 | 94.75 | 98.1667 | 97.3333 | 101.3333 | 100.6667 | 99.3333 | 102.3333 | 90.5544 | 99.8674 | 99.2 | 68.5917 | 92.7136 | 104.3 | 95.3333 | 100.267 | 98.5 | 98.1667 | 108.3333 | 99.6667 | 107.0833 | 93.1862 | 91.4832 | 86.5 | 99.3333 | 90.5
| Median |  | 97.5 | 97 | 94.5 | 98 | 97.5 | 101.5 | 100 | 99.5 | 102 | 90.4593 | 99.2164 | 98.8 | 67.975 | 93.368 | 104.35 | 95.5 | 100.4572 | 99 | 99 | 108.5 | 100 | 107.1 | 93.15 | 91.8159 | 87 | 99.5 | 90.5
| Std Dev. |  | 1.7512 | 1.6432 | 0.90719 | 0.75277 | 1.9664 | 1.2111 | 2.0656 | 1.633 | 1.5055 | 0.65532 | 1.4321 | 1.5937 | 3.1635 | 6.044 | 2.3281 | 0.8165 | 0.95743 | 0.83666 | 1.3292 | 2.582 | 1.3663 | 1.1321 | 0.90795 | 1.1957 | 3.5637 | 4.3665 | 1.6432
| Rel.Std.Dev. |  | 1.799% | 1.703% | 0.9575% | 0.7668% | 2.02% | 1.195% | 2.052% | 1.644% | 1.471% | 0.7237% | 1.434% | 1.607% | 4.612% | 6.519% | 2.232% | 0.8565% | 0.9549% | 0.8494% | 1.354% | 2.383% | 1.371% | 1.057% | 0.9743% | 1.307% | 4.12% | 4.396% | 1.816%
| PDM3 |  | -0.95% | -1.798% | -3.579% | -0.102% | -0.95% | 3.121% | 2.442% | 1.085% | 4.138% | -7.848% | 1.629% | 0.9496% | -30.2% | -5.651% | 6.14% | -2.985% | 2.035% | 0.2372% | -0.102% | 10.24% | 1.424% | 8.972% | -5.17% | -6.903% | -11.97% | 1.085% | -7.904%
| Z-Score (Absolute) |  | 0.2 | 0.4 | 0.81 | 0 | 0.2 | 0.75 | 0.59 | 0.28 | 0.99 | 1.81 | 0.4 | 0.25 | 7.04 | 1.3 | 1.46 | 0.67 | 0.5 | 0.08 | 0 | 2.42 | 0.36 | 2.12 | 1.19 | 1.59 | 2.78 | 0.28 | 1.82
| Analytical results for Zr in OREAS 505b (Certified Value 68 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES |  | 0
1
| 1 | 1 | 64.6 | 64.9 | 58.2 | 69.4 | 67.5 | 63.8 | 63.1 | 71.4 | 71.7 | 72.2277 | 67.3788 | 73.3 | 62.446 | 75.04 | 74.6 | 70.9389 | 66 | 69.2 | 67.6 | 67.6 | 65 | 68.14 | 48.6743 | 65.3 | 85.9 | 59.3 |  | 1
| 1 | 2 | 63.9 | 61.7 | 60.3 | 67.9 | 69.2 | 67.2 | 66.7 | 71.5 | 68.6 | 69.0227 | 67.9191 | 72.1 | 60.597 | 76.79 | 74.3 | 73.1835 | 66 | 67.6 | 67.8 | 72.8 | 67.2 | 70.19 | 46.9941 | 65.5 | 82.6 | 55.7 |  | 19
| 1 | 3 | 61.6 | 63 | 59.1 | 66.7 | 70.6 | 69.3 | 67.2 | 65.5 | 68.8 | 69.1777 | 67.7744 | 71.7 | 65.097 | 75.19 | 72.8 | 70.9013 | 68 | 66.4 | 64.2 | 71.7 | 64.9 | 68.91 | 49.2165 | 65.7 | 87.2 | 54.6 |  | 16
| 1 | 4 | 64.9 | 65.6 | 58.4 | 65.6 | 72.5 | 65.9 | 67.8 | 71.3 | 73.9 | 75.1549 | 67.5411 | 69.5 | 63.936 | 74.57 | 72.7 | 70.7402 | 66 | 67.7 | 68.9 | 72.8 | 76.3 | 70.13 | 50.254 | 65 | 85.8 | 54.7 |  | 68.2273
| 1 | 5 | 64.3 | 61.6 | 58.4 | 67.9 | 72 | 68.9 | 65.1 | 70.9 | 74.6 | 66.6678 | 67.1468 | 71.3 | 64.926 | 75.51 | 72 | 73.9825 | 67 | 69 | 71.5 | 71.1 | 66.5 | 71.2 | 48.1517 | 63.4 | 81.9 | 56.5 |  | 69
| 1 | 6 | 64.3 | 65.6 | 60 | 64.5 | 69.6 | 67 | 67 | 66.4 | 72.6 | 67.7256 | 67.7432 | 69.9 | 63.459 | 75.56 | 73.7 | 73.9313 | 65 | 69.2 | 67.3 | 72.9 | 70.4 | 71.1 | 49.0102 | 65.4 | 81.1 | 58.1
| Mean |  | 63.9333 | 63.7333 | 59.0667 | 67 | 70.2333 | 67.0167 | 66.15 | 69.5 | 71.7 | 69.9961 | 67.5839 | 71.3 | 63.4102 | 75.4433 | 73.35 | 72.2796 | 66.3333 | 68.1833 | 67.8833 | 71.4833 | 68.3833 | 69.945 | 48.7168 | 65.05 | 84.0833 | 56.4833
| Median |  | 64.3 | 63.95 | 58.75 | 67.3 | 70.1 | 67.1 | 66.85 | 71.1 | 72.15 | 69.1002 | 67.6421 | 71.5 | 63.6975 | 75.35 | 73.25 | 72.0612 | 66 | 68.35 | 67.7 | 72.25 | 66.85 | 70.16 | 48.8423 | 65.35 | 84.2 | 56.1
| Std Dev. |  | 1.1911 | 1.8737 | 0.89815 | 1.7709 | 1.8619 | 2.0193 | 1.7467 | 2.772 | 2.533 | 3.1446 | 0.28601 | 1.4142 | 1.6891 | 0.75115 | 1.0134 | 1.5819 | 1.0328 | 1.1392 | 2.371 | 2.0371 | 4.3651 | 1.211 | 1.0943 | 0.84083 | 2.523 | 1.8915
| Rel.Std.Dev. |  | 1.863% | 2.94% | 1.521% | 2.643% | 2.651% | 3.013% | 2.641% | 3.988% | 3.533% | 4.493% | 0.4232% | 1.983% | 2.664% | 0.9956% | 1.382% | 2.189% | 1.557% | 1.671% | 3.493% | 2.85% | 6.383% | 1.731% | 2.246% | 1.293% | 3.001% | 3.349%
| PDM3 |  | -6.294% | -6.587% | -13.43% | -1.799% | 2.94% | -1.774% | -3.045% | 1.865% | 5.09% | 2.592% | -0.943% | 4.504% | -7.06% | 10.58% | 7.508% | 5.939% | -2.776% | -0.06445% | -0.5042% | 4.772% | 0.2287% | 2.518% | -28.6% | -4.657% | 23.24% | -17.21%
| Z-Score (Absolute) |  | 0.88 | 0.93 | 1.93 | 0.22 | 0.47 | 0.22 | 0.41 | 0.32 | 0.79 | 0.42 | 0.1 | 0.7 | 1 | 1.6 | 1.15 | 0.92 | 0.37 | 0.03 | 0.03 | 0.74 | 0.08 | 0.41 | 4.17 | 0.64 | 3.46 | 2.49
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C258: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C276: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C294: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C313: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C332: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C350: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C369: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C388: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C406: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C424: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C442: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C479: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C497: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C515: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C533: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C552: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C570: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C588: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C606: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C624: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C642: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C660: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C678: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C696: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C714: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C732: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C750: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C768: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C787: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C825: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C843: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C861: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C879: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C897: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C915: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C934: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C972: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C990: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1008: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1027: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1046: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1064: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1082: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1101: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1119: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1137: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1155: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 505b (Certified Value 1.28 ppm) | [28 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g |  |  | 3
| 1 | 1 | 1.26 | 1.2 | 1.24 | 1.41 | 1.22 | 1.33 | 1.4 | 1.32 | 1.18 | 1.2617 | 1.73 | 1.67 | 1.2 | 1.23 | 1.3611 | 1.31 | 1.22 | 1.32 | 1.29 | 1.253 | < 5 | 1.31 | 1.3 | 1.29 | < 2 |  |  | 1
| 1 | 2 | 1.27 | 1.2 | 1.241 | 1.38 | 1.24 | 1.35 | 1.33 | 1.25 | 1.16 | 1.3248 | 1.55 | 1.94 | 1.22 | 1.21 | 1.2946 | 1.39 | 1.23 | 1.34 | 1.25 | 1.245 | < 5 | 1.249 | 1.3 | 1.36 | < 2 |  |  | 25
| 1 | 3 | 1.26 | 1.3 | 1.246 | 1.38 | 1.24 | 1.39 | 1.38 | 1.28 | 1.2 | 1.3066 | 1.47 | 2.45 | 1.18 | 1.21 | 1.3498 | 1.38 | 1.25 | 1.33 | 1.25 | 1.248 | < 5 | 1.312 | 1.2 | 1.3 | < 2 |  |  | 16
| 1 | 4 | 1.25 | 1.2 | 1.218 | 1.39 | 1.2 | 1.36 | 1.38 | 1.27 | 1.16 | 1.2808 | 1.29 | 2.04 | 1.19 | 1.21 | 1.2859 | 1.33 | 1.22 | 1.26 | 1.22 | 1.246 | < 5 | 1.282 | 1.3 | 1.34 | < 2 |  |  | 1.2838
| 1 | 5 | 1.27 | 1.3 | 1.216 | 1.35 | 1.24 | 1.36 | 1.35 | 1.32 | 1.22 | 1.2595 | 1.26 | 1.286 | 1.23 | 1.24 | 1.3574 | 1.38 | 1.3 | 1.27 | 1.3 | 1.25 | < 5 | 1.306 | 1.2 | 1.29 | < 2 |  |  | 71
| 1 | 6 | 1.25 | 1.3 | 1.246 | 1.33 | 1.22 | 1.32 | 1.31 | 1.19 | 1.18 | 1.2787 | 1.3 | 1.4268 | 1.25 | 1.18 | 1.3459 | 1.42 | 1.27 | 1.31 | 1.28 | 1.261 | < 5 | 1.289 | 1.3 | 1.31 | < 2
| Mean |  | 1.26 | 1.25 | 1.2345 | 1.3733 | 1.2267 | 1.3517 | 1.3583 | 1.2717 | 1.1833 | 1.2853 | 1.4333 | 1.8021 | 1.2117 | 1.2133 | 1.3325 | 1.3683 | 1.2483 | 1.305 | 1.265 | 1.2505 |  | 1.2913 | 1.2667 | 1.315
| Median |  | 1.26 | 1.25 | 1.2405 | 1.38 | 1.23 | 1.355 | 1.365 | 1.275 | 1.18 | 1.2797 | 1.385 | 1.805 | 1.21 | 1.21 | 1.3479 | 1.38 | 1.24 | 1.315 | 1.265 | 1.249 |  | 1.2975 | 1.3 | 1.305
| Std Dev. |  | 0.0089443 | 0.054772 | 0.013795 | 0.028752 | 0.01633 | 0.024833 | 0.034303 | 0.048751 | 0.023381 | 0.025693 | 0.18511 | 0.42896 | 0.026394 | 0.020656 | 0.033245 | 0.040702 | 0.031885 | 0.032711 | 0.030166 | 0.0058907 |  | 0.023981 | 0.05164 | 0.02881
| Rel.Std.Dev. |  | 0.7099% | 4.382% | 1.117% | 2.094% | 1.331% | 1.837% | 2.525% | 3.834% | 1.976% | 1.999% | 12.91% | 23.8% | 2.178% | 1.702% | 2.495% | 2.975% | 2.554% | 2.507% | 2.385% | 0.4711% |  | 1.857% | 4.077% | 2.191%
| PDM3 |  | -1.853% | -2.632% | -3.839% | 6.975% | -4.45% | 5.287% | 5.807% | -0.9443% | -7.825% | 0.1206% | 11.65% | 40.38% | -5.618% | -5.488% | 3.79% | 6.585% | -2.762% | 1.652% | -1.464% | -2.593% |  | 0.5876% | -1.334% | 2.431%
| Z-Score (Absolute) |  | 0.29 | NA | 0.58 | 1.01 | 0.67 | 0.76 | 0.84 | 0.16 | 1.17 | 0 | 1.7 | 5.94 | 0.85 | 0.83 | 0.54 | 0.95 | 0.43 | 0.23 | 0.23 | 0.4 | 13.96 | 0.07 | NA | 0.34 | 3.28
Results from laboratories 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Al in OREAS 505b (Certified Value 1.9 wt.%) | [28 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 3
| 1 | 1 | 1.93 | 1.56 | 1.86 | 1.88 | 2 | 1.98 | 1.99 | 1.86 | 1.751 | 1.8751 | 1.8929 | 1.64 | 2.03 | 1.8958 | 1.8 | 1.93 | 1.9239 | 1.84 | 1.88 | 2.11 | 1.96 | 1.53 | 1.9072 | 1.89 | 1.89 | 1.758 |  | 1
| 1 | 2 | 1.9 | 1.6 | 1.82 | 1.88 | 1.97 | 1.98 | 1.97 | 1.87 | 1.737 | 1.8369 | 1.8478 | 1.64 | 1.649 | 1.8996 | 1.83 | 1.92 | 1.9938 | 1.85 | 1.87 | 2.15 | 1.99 | 1.57 | 1.9338 | 1.89 | 1.94 | 1.78 |  | #N/A
| 1 | 3 | 1.93 | 1.65 | 1.85 | 1.88 | 1.98 | 1.98 | 1.98 | 1.86 | 1.71 | 1.8394 | 1.8469 | 1.65 | 1.95 | 1.9322 | 1.86 | 1.95 | 1.953 | 1.82 | 1.86 | 2.1 | 1.99 | 1.66 | 1.8871 | 1.91 | 1.92 | 1.76 |  | 16
| 1 | 4 | 1.94 | 1.66 | 1.84 | 1.91 | 1.98 | 1.97 | 1.98 | 1.88 | 1.679 | 1.8604 | 1.8629 | 1.63 | 1.065 | 1.9377 | 1.81 | 1.91 | 1.9048 | 1.84 | 1.86 | 1.98 | 1.99 | 1.61 | 1.8958 | 1.88 | 1.92 | 1.737 |  | 1.9006
| 1 | 5 | 1.96 | 1.68 | 1.84 | 1.92 | 1.97 | 1.96 | 1.94 | 1.88 | 1.729 | 1.8637 | 1.848 | 1.68 | 2.03 | 1.9337 | 1.78 | 1.96 | 1.9791 | 1.85 | 1.88 | 2 | 2 | 1.68 | 1.9309 | 1.86 | 1.91 | 1.755 |  | 72
| 1 | 6 | 1.96 | 1.63 | 1.85 | 1.92 | 1.94 | 1.96 | 1.93 | 1.91 | 1.759 | 1.8653 | 1.8406 | 1.62 | 1.83 | 1.9486 | 1.83 | 1.9 | 1.9396 | 1.83 | 1.82 | 2.07 | 2.02 | 1.66 | 1.9547 | 1.87 | 1.96 | 1.758
| Mean |  | 1.9367 | 1.63 | 1.8433 | 1.8983 | 1.9733 | 1.9717 | 1.965 | 1.8767 | 1.7275 | 1.8568 | 1.8565 | 1.6433 | 1.759 | 1.9246 | 1.8183 | 1.9283 | 1.949 | 1.8383 | 1.8617 | 2.0683 | 1.9917 | 1.6183 | 1.9183 | 1.8833 | 1.9233 | 1.758
| Median |  | 1.935 | 1.64 | 1.845 | 1.895 | 1.975 | 1.975 | 1.975 | 1.875 | 1.733 | 1.862 | 1.8479 | 1.64 | 1.89 | 1.9329 | 1.82 | 1.925 | 1.9463 | 1.84 | 1.865 | 2.085 | 1.99 | 1.635 | 1.9191 | 1.885 | 1.92 | 1.758
| Std Dev. |  | 0.022509 | 0.043818 | 0.013663 | 0.020412 | 0.019664 | 0.0098319 | 0.02429 | 0.018619 | 0.02931 | 0.015251 | 0.019288 | 0.020656 | 0.36936 | 0.021645 | 0.027869 | 0.023166 | 0.033489 | 0.01169 | 0.022286 | 0.066156 | 0.019408 | 0.059133 | 0.025787 | 0.017512 | 0.024221 | 0.013697
| Rel.Std.Dev. |  | 1.162% | 2.688% | 0.7412% | 1.075% | 0.9965% | 0.4987% | 1.236% | 0.9921% | 1.697% | 0.8214% | 1.039% | 1.257% | 21% | 1.125% | 1.533% | 1.201% | 1.718% | 0.6359% | 1.197% | 3.199% | 0.9745% | 3.654% | 1.344% | 0.9298% | 1.259% | 0.7791%
| PDM3 |  | 1.9% | -14.24% | -3.011% | -0.1167% | 3.83% | 3.742% | 3.391% | -1.257% | -9.105% | -2.303% | -2.318% | -13.53% | -7.448% | 1.265% | -4.326% | 1.462% | 2.551% | -3.274% | -2.046% | 8.828% | 4.794% | -14.85% | 0.9319% | -0.906% | 1.199% | -7.501%
| Z-Score (Absolute) |  | 0.65 | 2.85 | 0.42 | 0.21 | 1.06 | 1.04 | 0.97 | 0.04 | 1.74 | 0.26 | 0.27 | 2.7 | 1.38 | 0.51 | 0.7 | 0.55 | 0.79 | 0.47 | 0.21 | 2.15 | 1.27 | 2.98 | 0.44 | 0.04 | 0.49 | 1.39
| Analytical results for As in OREAS 505b (Certified Value 9.01 ppm) | [28 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g |  | 2
| 1 | 1 | 9 | 7.7 | 8.6 | 9 | 9.3 | 9.6 | 9.3 | 8.6 | 8.87 | 6.8196 | 9.9393 | 27.5 | 5.09 | 10.79 | 6.7 | 10 | 7.7582 | 9.6 | 8.7 | 9.1 | 9.8 | 7.261 | 7.889 | 8.7 | 10 | 13 |  | 1
| 1 | 2 | 10 | 8.1 | 8.2 | 9 | 9.3 | 9.6 | 8.8 | 8.9 | 9.08 | 6.1202 | 9.4093 | 10 | 6.28 | 10.08 | 6.6 | 9 | 8.9224 | 9.5 | 8.3 | 9.7 | 9.7 | 7.577 | 8.756 | 9.4 | 10 | 12 |  | 13
| 1 | 3 | 9 | 8.5 | 8.2 | 9 | 9.2 | 9.2 | 8.8 | 8.9 | 9.22 | 8.9525 | 10.2813 | 9.9 | 5.8 | 10.68 | 6 | 10 | 8.5978 | 9.3 | 8.7 | 9.6 | 9.5 | 6.69 | 8.521 | 8.6 | 11 | 15 |  | 16
| 1 | 4 | 9 | 8.2 | 8 | 9 | 9.2 | 9.3 | 9.9 | 8.7 | 9.21 | 7.4835 | 9.7908 | 11.1 | 6.06 | 10.48 | 6.6 | 9 | 7.9465 | 9.5 | 8.5 | 9.4 | 9.6 | 7.156 | 8.163 | 8.8 | 10 | 13 |  | 9.0064
| 1 | 5 | 9 | 8.4 | 7.9 | 8 | 9.7 | 9.7 | 9.3 | 9.1 | 9.39 | 7.5512 | 9.2879 | 9.7 | 4.91 | 10.71 | 6.9 | 11 | 8.6718 | 9.4 | 9.7 | 8.9 | 9.4 | 6.637 | 8.806 | 8.6 | 10 | 9 |  | 73
| 1 | 6 | 9 | 8.2 | 8 | 8 | 9.4 | 9.5 | 9.3 | 8.7 | 9.29 | 7.2287 | 9.4241 | 10.1 | 5.6 | 10.54 | 6.5 | 10 | 8.9354 | 9.7 | 9.1 | 9 | 10.2 | 6.797 | 8.863 | 8.6 | 11 | 11
| Mean |  | 9.1667 | 8.1833 | 8.15 | 8.6667 | 9.35 | 9.4833 | 9.2333 | 8.8167 | 9.1767 | 7.3593 | 9.6888 | 13.05 | 5.6233 | 10.5467 | 6.55 | 9.8333 | 8.472 | 9.5 | 8.8333 | 9.2833 | 9.7 | 7.0197 | 8.4997 | 8.7833 | 10.3333 | 12.1667
| Median |  | 9 | 8.2 | 8.1 | 9 | 9.3 | 9.55 | 9.3 | 8.8 | 9.215 | 7.3561 | 9.6075 | 10.05 | 5.7 | 10.61 | 6.6 | 10 | 8.6348 | 9.5 | 8.7 | 9.25 | 9.65 | 6.9765 | 8.6385 | 8.65 | 10 | 12.5
| Std Dev. |  | 0.40825 | 0.27869 | 0.251 | 0.5164 | 0.18708 | 0.19408 | 0.40825 | 0.18348 | 0.18151 | 0.94173 | 0.38292 | 7.0958 | 0.53795 | 0.25532 | 0.30166 | 0.75277 | 0.50178 | 0.14142 | 0.50067 | 0.33116 | 0.28284 | 0.37206 | 0.39453 | 0.31252 | 0.5164 | 2.0412
| Rel.Std.Dev. |  | 4.454% | 3.406% | 3.08% | 5.958% | 2.001% | 2.047% | 4.421% | 2.081% | 1.978% | 12.8% | 3.952% | 54.37% | 9.566% | 2.421% | 4.606% | 7.655% | 5.923% | 1.489% | 5.668% | 3.567% | 2.916% | 5.3% | 4.642% | 3.558% | 4.997% | 16.78%
| PDM3 |  | 1.78% | -9.138% | -9.508% | -3.772% | 3.815% | 5.296% | 2.52% | -2.106% | 1.891% | -18.29% | 7.577% | 44.9% | -37.56% | 17.1% | -27.27% | 9.182% | -5.933% | 5.481% | -1.921% | 3.075% | 7.702% | -22.06% | -5.626% | -2.476% | 14.73% | 35.09%
| Z-Score (Absolute) |  | NA | 0.91 | 0.95 | NA | 0.38 | 0.53 | 0.25 | 0.21 | 0.19 | 1.82 | 0.76 | 4.48 | 3.75 | 1.71 | 2.72 | 0.92 | 0.59 | 0.55 | 0.19 | 0.31 | 0.77 | 2.2 | 0.56 | 0.25 | NA | NA
Results from laboratories 1.01, 1.04, 1.29 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for B in OREAS 505b (Certified Value < 10 ppm) | [28 empty] | Certified
B | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.03; 1.03: 1.05; 1.04: 1.06; 1.05: 1.07; 1.06: 1.08; 1.07: 1.10; 1.08: 1.12; 1.09: 1.13; 1.10: 1.18; 1.11: 1.20; 1.12: 1.21; 1.13: 1.23; 1.14: 1.29 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES; 1.08: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*OES/MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS |  | 1
|  |  | 1.01: 0.5g; 1.02: 15g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.5g; 1.09: 0.5g; 1.10: 01g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g; 1.14: 0.25g |  | 1
| 1 | 1 | 1.01: < 20; 1.02: < 1; 1.03: < 10; 1.04: < 10; 1.05: < 10; 1.06: < 10; 1.07: 9.6926; 1.08: 26; 1.09: 19.38; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10 |  | 1
| 1 | 2 | 1.01: < 20; 1.02: < 1; 1.03: < 10; 1.04: < 10; 1.05: < 10; 1.06: < 10; 1.07: 10.0607; 1.08: 26; 1.09: 14.5; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10 |  | 3
| 1 | 3 | 1.01: < 20; 1.02: < 1; 1.03: < 10; 1.04: < 10; 1.05: < 10; 1.06: < 10; 1.07: 8.9549; 1.08: 27; 1.09: 22.71; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10 |  | 16
| 1 | 4 | 1.01: < 20; 1.02: < 1; 1.03: < 10; 1.04: < 10; 1.05: 10; 1.06: < 10; 1.07: 9.4014; 1.08: 27; 1.09: 24.95; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10 |  | < 10
| 1 | 5 | 1.01: < 20; 1.02: < 1; 1.03: < 10; 1.04: < 10; 1.05: 10; 1.06: < 10; 1.07: 11.6145; 1.08: 26; 1.09: 22.85; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10 |  | 74
| 1 | 6 | 1.01: < 20; 1.02: < 1; 1.03: < 10; 1.04: < 10; 1.05: 10; 1.06: < 10; 1.07: 10.1405; 1.08: 27; 1.09: 27.26; 1.10: < 10; 1.11: < 10; 1.12: < 10; 1.13: < 20; 1.14: < 10
| Mean |  | 1.05: 10; 1.07: 9.9774; 1.08: 26.5; 1.09: 21.9417
| Median |  | 1.05: 10; 1.07: 9.8766; 1.08: 26.5; 1.09: 22.78
| Std Dev. |  | 1.05: 0; 1.07: 0.91346; 1.08: 0.54772; 1.09: 4.4876
| Rel.Std.Dev. |  | 1.05: 0%; 1.07: 9.155%; 1.08: 2.067%; 1.09: 20.45%
PDM3
| Z-Score (Absolute) |  | 1.01: 2.7; 1.02: 2.43; 1.03: 0; 1.04: 0; 1.05: 1.35; 1.06: 0; 1.07: 2.69; 1.08: 11.6; 1.09: 9.14; 1.10: 0; 1.11: 0; 1.12: 0; 1.13: 2.7; 1.14: 0
| Analytical results for Ba in OREAS 505b (Certified Value 454 ppm) | [28 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  | 0
| 1 | 1 | 457 | 420 | 384.9 | 483 | 490 | 450 | 470 | 470 | 251.35 | 421.8165 | 491.8319 | 467 | 484.52 | 476.7 | 511 | 417.6213 | 474 | 470 | 420 | 310 | 409.9 | 435.388 | 431 | 562 | 443 |  |  | 1
| 1 | 2 | 456 | 430 | 350.8 | 483 | 490 | 450 | 470 | 470 | 262.73 | 409.8945 | 492.3689 | 473 | 473.33 | 480.4 | 515 | 431.9425 | 438 | 470 | 400 | 308 | 416.1 | 436.34 | 433 | 596 | 443 |  |  | 14
| 1 | 3 | 449 | 439 | 386.6 | 485 | 490 | 450 | 480 | 470 | 272.6 | 416.4909 | 482.3841 | 468 | 480.03 | 473.7 | 508 | 398.5643 | 470 | 460 | 430 | 305 | 435.7 | 439.467 | 434 | 571 | 440 |  |  | 16
| 1 | 4 | 453 | 441 | 390.7 | 488 | 490 | 450 | 480 | 480 | 254.75 | 418.4861 | 488.089 | 464 | 480.87 | 469.1 | 505 | 425.266 | 463 | 460 | 390 | 309 | 426.4 | 436.616 | 428 | 586 | 439 |  |  | 453.678
| 1 | 5 | 451 | 444 | 372.8 | 484 | 490 | 440 | 480 | 450 | 298.98 | 420.4959 | 488.2985 | 474 | 472.54 | 483.2 | 516 | 438.821 | 465 | 470 | 220 | 311 | 436.7 | 445.144 | 423 | 574 | 443 |  |  | 75
| 1 | 6 | 459 | 437 | 381.2 | 485 | 480 | 450 | 480 | 470 | 288.75 | 418.6399 | 487.7746 | 463 | 471.1 | 479.4 | 504 | 401.0953 | 477 | 460 | 430 | 310 | 434.7 | 442.031 | 420 | 601 | 445
| Mean |  | 454.1667 | 435.1667 | 377.8333 | 484.6667 | 488.3333 | 448.3333 | 476.6667 | 468.3333 | 271.5267 | 417.6373 | 488.4578 | 468.1667 | 477.065 | 477.0833 | 509.8333 | 418.8851 | 464.5 | 465 | 381.6667 | 308.8333 | 426.5833 | 439.1643 | 428.1667 | 581.6667 | 442.1667
| Median |  | 454.5 | 438 | 383.05 | 484.5 | 490 | 450 | 480 | 470 | 267.665 | 418.563 | 488.1937 | 467.5 | 476.68 | 478.05 | 509.5 | 421.4437 | 467.5 | 465 | 410 | 309.5 | 430.55 | 438.0415 | 429.5 | 580 | 443
| Std Dev. |  | 3.8166 | 8.7958 | 14.5589 | 1.8619 | 4.0825 | 4.0825 | 5.164 | 9.8319 | 19.0652 | 4.2102 | 3.5833 | 4.535 | 5.4563 | 5.0815 | 5.0365 | 16.3686 | 14.0107 | 5.4772 | 80.8497 | 2.137 | 11.3102 | 3.8157 | 5.6362 | 15.214 | 2.2286
| Rel.Std.Dev. |  | 0.8404% | 2.021% | 3.853% | 0.3842% | 0.836% | 0.9106% | 1.083% | 2.099% | 7.021% | 1.008% | 0.7336% | 0.9687% | 1.144% | 1.065% | 0.9879% | 3.908% | 3.016% | 1.178% | 21.18% | 0.692% | 2.651% | 0.8689% | 1.316% | 2.616% | 0.504%
| PDM3 |  | 0.1077% | -4.08% | -16.72% | 6.831% | 7.639% | -1.178% | 5.067% | 3.23% | -40.15% | -7.944% | 7.666% | 3.194% | 5.155% | 5.159% | 12.38% | -7.669% | 2.385% | 2.496% | -15.87% | -31.93% | -5.972% | -3.199% | -5.623% | 28.21% | -2.537%
| Z-Score (Absolute) |  | 0 | 0.49 | 1.98 | 0.79 | 0.89 | 0.15 | 0.58 | 0.37 | 4.74 | 0.95 | 0.89 | 0.36 | 0.59 | 0.59 | 1.44 | 0.92 | 0.27 | 0.28 | 1.88 | 3.77 | 0.72 | 0.39 | 0.67 | 3.31 | 0.31
| Analytical results for Be in OREAS 505b (Certified Value 1.48 ppm) | [28 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | [6 empty] | 3
| 1 | 1 | 1.4 | 1.6 | 1.45 | 1.42 | 1.46 | 1.38 | 1.57 | 0.9851 | 1.4304 | 2.33 | < 2 | 1.9 | 1.2829 | 1.56 | 1.35 | 1.53 | < 5 | < 5 | 2 | 1.7 | 1.6 | [6 empty] | 1
| 1 | 2 | 1.4 | 1.6 | 1.45 | 1.4 | 1.45 | 1.39 | 1.55 | 0.9733 | 1.394 | 2.36 | < 2 | 1.9 | 1.3832 | 1.53 | 1.37 | 1.59 | < 5 | < 5 | 1 | 1.7 | 1.6 | [6 empty] | 28
| 1 | 3 | 1.3 | 1.6 | 1.46 | 1.41 | 1.45 | 1.39 | 1.56 | 1.007 | 1.3881 | 2.4 | < 2 | 2 | 1.3579 | 1.6 | 1.35 | 1.53 | < 5 | < 5 | 2 | 1.7 | 1.6 | [6 empty] | 16
| 1 | 4 | 1.5 | 1.6 | 1.45 | 1.4 | 1.5 | 1.4 | 1.58 | 0.9333 | 1.4007 | 2.33 | < 2 | 1.9 | 1.4016 | 1.56 | 1.35 | 1.46 | < 5 | < 5 | 2 | 1.7 | 1.6 | [6 empty] | 1.4834
| 1 | 5 | 1.4 | 1.7 | 1.44 | 1.38 | 1.48 | 1.39 | 1.57 | 0.9689 | 1.4007 | 2.28 | < 2 | 1.9 | 1.4108 | 1.58 | 1.37 | 1.46 | < 5 | < 5 | 1 | 1.8 | 1.6 | [6 empty] | 76
| 1 | 6 | 1.3 | 1.6 | 1.42 | 1.38 | 1.47 | 1.41 | 1.6 | 0.9529 | 1.3759 | 2.36 | < 2 | 1.8 | 1.2753 | 1.53 | 1.32 | 1.5 | < 5 | < 5 | 1 | 1.7 | 1.6
| Mean |  | 1.3833 | 1.6167 | 1.445 | 1.3983 | 1.4683 | 1.3933 | 1.5717 | 0.97008 | 1.3983 | 2.3433 |  | 1.9 | 1.3519 | 1.56 | 1.3517 | 1.5117 |  |  | 1.5 | 1.7167 | 1.6
| Median |  | 1.4 | 1.6 | 1.45 | 1.4 | 1.465 | 1.39 | 1.57 | 0.9711 | 1.3973 | 2.345 |  | 1.9 | 1.3706 | 1.56 | 1.35 | 1.515 |  |  | 1.5 | 1.7 | 1.6
| Std Dev. |  | 0.075277 | 0.040825 | 0.013784 | 0.016021 | 0.019408 | 0.010328 | 0.017224 | 0.025489 | 0.018243 | 0.040332 |  | 0.063246 | 0.059313 | 0.027568 | 0.018348 | 0.049565 |  |  | 0.54772 | 0.040825 | 2.4324e-16
| Rel.Std.Dev. |  | 5.442% | 2.525% | 0.9539% | 1.146% | 1.322% | 0.7412% | 1.096% | 2.627% | 1.305% | 1.721% |  | 3.329% | 4.387% | 1.767% | 1.357% | 3.279% |  |  | 36.51% | 2.378% | 1.52e-14%
| PDM3 |  | -6.743% | 8.987% | -2.585% | -5.731% | -1.012% | -6.069% | 5.954% | -34.6% | -5.733% | 57.98% |  | 28.09% | -8.859% | 5.167% | -8.877% | 1.909% |  |  | 1.122% | 15.73% | 7.864%
| Z-Score (Absolute) |  | 0.61 | 0.72 | 0.26 | 0.52 | 0.12 | 0.55 | 0.47 | 2.96 | 0.52 | 4.86 | 2.79 | 2.34 | 0.79 | 0.4 | 0.79 | 0.12 | 5.76 | 5.76 | NA | 1.29 | 0.63
Results from laboratory 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 505b (Certified Value 0.75 ppm) | [28 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g |  | 2
| 1 | 1 | 0.75 | 0.8 | 0.73 | 0.65 | 0.61 | 0.6 | 0.93 | 0.7 | 0.72 | < 0.1 | 0.89556 | 0.83 | 0.49 | < 5 | 1.04 | 0.86555 | 0.53 | 0.61 | 1.2 | < 2 | 0.93 | < 5 | 0.837 | 0.7 | 0.74 | < 5 |  | 1
| 1 | 2 | 0.74 | 0.7 | 0.72 | 0.78 | 0.6 | 0.74 | 0.67 | 0.87 | 0.75 | < 0.1 | 0.77642 | 0.76 | 0.93 | < 5 | 1.23 | 0.75992 | 0.51 | 0.6 | 0.86 | < 2 | 0.78 | < 5 | 0.575 | 0.7 | 0.82 | < 5 |  | 29
| 1 | 3 | 0.74 | 1 | 0.7 | 0.72 | 0.71 | 0.68 | 0.63 | 0.78 | 0.7 | < 0.1 | 0.75575 | 0.76 | 1.46 | < 5 | 1.06 | 0.80153 | 0.43 | 0.64 | 1.06 | < 2 | 0.75 | < 5 | 0.614 | 0.7 | 0.91 | < 5 |  | 16
| 1 | 4 | 0.71 | 0.7 | 0.75 | 0.77 | 0.68 | 0.75 | 0.87 | 0.7 | 0.86 | 0.8935 | 0.83977 | 0.73 | 1.45 | < 5 | 0.97 | 0.69739 | 0.36 | 0.61 | 0.88 | < 2 | 0.96 | < 5 | 0.583 | 0.7 | 0.76 | < 5 |  | 0.74516
| 1 | 5 | 0.76 | 1.4 | 0.68 | 0.75 | 0.66 | 0.67 | 0.71 | 0.73 | 0.68 | < 0.1 | 0.68109 | 0.74 | 1.34 | < 5 | 1.25 | 0.86632 | 0.54 | 0.67 | 0.66 | < 2 | 0.72 | < 5 | 0.609 | 0.7 | 0.92 | < 5 |  | 77
| 1 | 6 | 0.76 | 0.7 | 0.7 | 0.63 | 0.73 | 0.7 | 0.68 | 0.78 | 0.66 | < 0.1 | 0.74059 | 0.93 | 1.22 | < 5 | 0.83 | 0.74313 | 0.4 | 0.59 | 0.73 | < 2 | 0.81 | < 5 | 0.601 | 0.6 | 0.78 | < 5
| Mean |  | 0.74333 | 0.88333 | 0.71333 | 0.71667 | 0.665 | 0.69 | 0.74833 | 0.76 | 0.72833 | 0.8935 | 0.78153 | 0.79167 | 1.1483 |  | 1.0633 | 0.78897 | 0.46167 | 0.62 | 0.89833 |  | 0.825 |  | 0.6365 | 0.68333 | 0.82167
| Median |  | 0.745 | 0.75 | 0.71 | 0.735 | 0.67 | 0.69 | 0.695 | 0.755 | 0.71 | 0.8935 | 0.76608 | 0.76 | 1.28 |  | 1.05 | 0.78073 | 0.47 | 0.61 | 0.87 |  | 0.795 |  | 0.605 | 0.7 | 0.8
| Std Dev. |  | 0.018619 | 0.27869 | 0.025033 | 0.06314 | 0.05244 | 0.054406 | 0.12172 | 0.064807 | 0.071671 |  | 0.075962 | 0.076267 | 0.37712 |  | 0.15895 | 0.068317 | 0.075211 | 0.029665 | 0.20203 |  | 0.098133 |  | 0.099365 | 0.040825 | 0.07705
| Rel.Std.Dev. |  | 2.505% | 31.55% | 3.509% | 8.81% | 7.886% | 7.885% | 16.27% | 8.527% | 9.84% |  | 9.72% | 9.634% | 32.84% |  | 14.95% | 8.659% | 16.29% | 4.785% | 22.49% |  | 11.89% |  | 15.61% | 5.974% | 9.377%
| PDM3 |  | -0.2453% | 18.54% | -4.271% | -3.824% | -10.76% | -7.403% | 0.4257% | 1.991% | -2.258% | 19.91% | 4.88% | 6.241% | 54.11% |  | 42.7% | 5.88% | -38.04% | -16.8% | 20.56% |  | 10.71% |  | -14.58% | -8.297% | 10.27%
| Z-Score (Absolute) |  | 0.2 | NA | 0.42 | 0.39 | 0.76 | 0.58 | 0.16 | 0.08 | 0.31 | 4.19 | 0.08 | 0.15 | 2.73 | 12.5 | 2.12 | 0.13 | 2.23 | 1.09 | 0.92 | 1.66 | 0.39 | 12.5 | 0.97 | NA | 0.37 | 12.5
Results from laboratories 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ca in OREAS 505b (Certified Value 0.785 wt.%) | [28 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 3
| 1 | 1 | 0.77 | 0.76 | 0.8 | 0.81 | 0.82 | 0.84 | 0.85 | 0.79 | 0.778 | 0.75434 | 0.79613 | 0.71 | 0.60513 | 0.87165 | 0.74 | 0.79 | 0.72168 | 0.79 | 0.79 | 0.87 | 0.85 | 0.7 | 0.78169 | 0.71614 | 0.79 | 0.77 | 0.724 | 1
| 1 | 2 | 0.77 | 0.78 | 0.79 | 0.81 | 0.84 | 0.83 | 0.84 | 0.8 | 0.783 | 0.7279 | 0.77548 | 0.72 | 0.54183 | 0.87223 | 0.76 | 0.8 | 0.73677 | 0.79 | 0.79 | 0.9 | 0.85 | 0.71 | 0.80185 | 0.69214 | 0.79 | 0.78 | 0.729 | #N/A
| 1 | 3 | 0.77 | 0.79 | 0.8 | 0.81 | 0.82 | 0.83 | 0.85 | 0.8 | 0.794 | 0.76246 | 0.76948 | 0.72 | 0.58371 | 0.87214 | 0.76 | 0.82 | 0.71737 | 0.78 | 0.79 | 0.86 | 0.86 | 0.72 | 0.80832 | 0.70035 | 0.8 | 0.77 | 0.721 | 16
| 1 | 4 | 0.77 | 0.79 | 0.79 | 0.82 | 0.82 | 0.83 | 0.81 | 0.8 | 0.775 | 0.7319 | 0.78486 | 0.72 | 0.64132 | 0.87168 | 0.74 | 0.78 | 0.71042 | 0.79 | 0.78 | 0.83 | 0.86 | 0.72 | 0.78698 | 0.71484 | 0.79 | 0.78 | 0.716 | 0.78532
| 1 | 5 | 0.78 | 0.8 | 0.79 | 0.82 | 0.84 | 0.83 | 0.79 | 0.8 | 0.809 | 0.7461 | 0.77884 | 0.73 | 0.63146 | 0.88661 | 0.75 | 0.82 | 0.73913 | 0.79 | 0.8 | 0.82 | 0.84 | 0.74 | 0.81551 | 0.70394 | 0.79 | 0.77 | 0.729 | 78
| 1 | 6 | 0.77 | 0.78 | 0.79 | 0.81 | 0.81 | 0.82 | 0.8 | 0.81 | 0.793 | 0.73896 | 0.77565 | 0.72 | 0.58397 | 0.8787 | 0.75 | 0.79 | 0.72898 | 0.78 | 0.77 | 0.85 | 0.87 | 0.74 | 0.8099 | 0.67986 | 0.77 | 0.78 | 0.728
| Mean |  | 0.77167 | 0.78333 | 0.79333 | 0.81333 | 0.825 | 0.83 | 0.82333 | 0.8 | 0.78867 | 0.74361 | 0.78007 | 0.72 | 0.5979 | 0.8755 | 0.75 | 0.8 | 0.72573 | 0.78667 | 0.78667 | 0.855 | 0.855 | 0.72167 | 0.80071 | 0.70121 | 0.78833 | 0.775 | 0.7245
| Median |  | 0.77 | 0.785 | 0.79 | 0.81 | 0.82 | 0.83 | 0.825 | 0.8 | 0.788 | 0.74253 | 0.77725 | 0.72 | 0.59455 | 0.87218 | 0.75 | 0.795 | 0.72533 | 0.79 | 0.79 | 0.855 | 0.855 | 0.72 | 0.80508 | 0.70215 | 0.79 | 0.775 | 0.726
| Std Dev. |  | 0.0040825 | 0.013663 | 0.005164 | 0.005164 | 0.012247 | 0.0063246 | 0.026583 | 0.0063246 | 0.012596 | 0.013285 | 0.0093298 | 0.0063246 | 0.036339 | 0.0060839 | 0.0089443 | 0.016733 | 0.011249 | 0.005164 | 0.010328 | 0.02881 | 0.010488 | 0.016021 | 0.013512 | 0.01382 | 0.0098319 | 0.0054772 | 0.005244
| Rel.Std.Dev. |  | 0.529% | 1.744% | 0.6509% | 0.6349% | 1.485% | 0.762% | 3.229% | 0.7906% | 1.597% | 1.787% | 1.196% | 0.8784% | 6.078% | 0.6949% | 1.193% | 2.092% | 1.55% | 0.6564% | 1.313% | 3.37% | 1.227% | 2.22% | 1.687% | 1.971% | 1.247% | 0.7067% | 0.7238%
| PDM3 |  | -1.739% | -0.253% | 1.02% | 3.567% | 5.053% | 5.689% | 4.84% | 1.869% | 0.4261% | -5.311% | -0.6682% | -8.318% | -23.87% | 11.48% | -4.498% | 1.869% | -7.589% | 0.1714% | 0.1714% | 8.873% | 8.873% | -8.105% | 1.959% | -10.71% | 0.3837% | -1.314% | -7.745%
| Z-Score (Absolute) |  | 0.28 | 0.06 | 0.12 | 0.49 | 0.7 | 0.8 | 0.67 | 0.25 | 0.04 | 0.79 | 0.12 | 1.23 | 3.47 | 1.63 | 0.67 | 0.25 | 1.12 | 0 | 0 | 1.26 | 1.26 | 1.2 | 0.26 | 1.57 | 0.03 | 0.21 | 1.14
| Analytical results for Cd in OREAS 505b (Certified Value 0.19 ppm) | [28 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  | 3
| 1 | 1 | 0.23 | 0.3 | 0.19 | 0.2 | 0.17 | 0.19 | 0.19 | 0.2 | 0.195 | < 0.1 | 0.2263 | 1.79 | 3.11 | < 1 | 0.26 | 0.20013 | 0.4 | 0.2 | 0.2 | < 0.5 | 0.21 | < 5 | 0.2 | 0.25 | < 1 |  |  | 1
| 1 | 2 | 0.22 | 0.3 | 0.18 | 0.2 | 0.18 | 0.16 | 0.18 | 0.18 | 0.206 | < 0.1 | 0.21088 | 0.32 | 3.12 | < 1 | 0.26 | 0.20428 | 0.38 | 0.18 | 0.19 | < 0.5 | 0.19 | < 5 | 0.2 | 0.28 | < 1 |  |  | 30
| 1 | 3 | 0.19 | 0.3 | 0.18 | 0.19 | 0.2 | 0.18 | 0.2 | 0.18 | 0.211 | < 0.1 | 0.2256 | 0.25 | 2.98 | < 1 | 0.25 | 0.18125 | 0.41 | 0.18 | 0.19 | < 0.5 | 0.2 | < 5 | 0.2 | 0.27 | < 1 |  |  | 16
| 1 | 4 | 0.21 | 0.3 | 0.18 | 0.19 | 0.17 | 0.19 | 0.21 | 0.18 | 0.204 | < 0.1 | 0.22637 | 0.22 | 3.17 | < 1 | 0.24 | 0.19692 | 0.4 | 0.18 | 0.18 | < 0.5 | 0.19 | < 5 | 0.2 | 0.27 | < 1 |  |  | 0.19382
| 1 | 5 | 0.18 | 0.3 | 0.19 | 0.2 | 0.18 | 0.18 | 0.2 | 0.19 | 0.205 | < 0.1 | 0.20019 | 0.22 | 3.03 | < 1 | 0.25 | 0.18159 | 0.4 | 0.2 | 0.2 | < 0.5 | 0.19 | < 5 | 0.2 | 0.26 | < 1 |  |  | 79
| 1 | 6 | 0.2 | 0.3 | 0.18 | 0.2 | 0.18 | 0.18 | 0.21 | 0.19 | 0.195 | < 0.1 | 0.21974 | 0.23 | 2.91 | < 1 | 0.24 | 0.19836 | 0.4 | 0.19 | 0.2 | < 0.5 | 0.18 | < 5 | 0.2 | 0.27 | < 1
| Mean |  | 0.205 | 0.3 | 0.18333 | 0.19667 | 0.18 | 0.18 | 0.19833 | 0.18667 | 0.20267 |  | 0.21818 | 0.505 | 3.0533 |  | 0.25 | 0.19376 | 0.39833 | 0.18833 | 0.19333 |  | 0.19333 |  | 0.2 | 0.26667
| Median |  | 0.205 | 0.3 | 0.18 | 0.2 | 0.18 | 0.18 | 0.2 | 0.185 | 0.2045 |  | 0.22267 | 0.24 | 3.07 |  | 0.25 | 0.19764 | 0.4 | 0.185 | 0.195 |  | 0.19 |  | 0.2 | 0.27
| Std Dev. |  | 0.018708 | 0 | 0.005164 | 0.005164 | 0.010954 | 0.010954 | 0.01169 | 0.008165 | 0.0064083 |  | 0.010651 | 0.63064 | 0.097707 |  | 0.0089443 | 0.0098674 | 0.0098319 | 0.0098319 | 0.008165 |  | 0.010328 |  | 3.0405e-17 | 0.010328
| Rel.Std.Dev. |  | 9.126% | 0% | 2.817% | 2.626% | 6.086% | 6.086% | 5.894% | 4.374% | 3.162% |  | 4.882% | 124.9% | 3.2% |  | 3.578% | 5.093% | 2.468% | 5.22% | 4.223% |  | 5.342% |  | 1.52e-14% | 3.873%
| PDM3 |  | 5.771% | 54.79% | -5.408% | 1.471% | -7.128% | -7.128% | 2.331% | -3.688% | 4.567% |  | 12.57% | 160.6% | 1475% |  | 28.99% | -0.03108% | 105.5% | -2.829% | -0.2488% |  | -0.2488% |  | 3.191% | 37.59%
| Z-Score (Absolute) |  | 0.07 | NA | 0.58 | 0.18 | 0.67 | 0.67 | 0.13 | 0.48 | 0 | 4.54 | 0.46 | 8.99 | 84.8 | 8.85 | 1.41 | 0.27 | 5.82 | 0.43 | 0.28 | 1.41 | 0.28 | 68.34 | NA | 1.9 | 8.85
Results from laboratories 1.02 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Ce in OREAS 505b (Certified Value 34 ppm) | [28 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.24 | 1.27 | 1.29 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | [9 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [9 empty] | 2
| 1 | 1 | 49.1 | 31.49 | 35.5 | 33.3 | 39.7 | 33.7 | 31.45 | 28.8595 | 39.0639 | 33.1 | 28 | 45.2795 | 35 | 34.8 | 34.9 | 41.203 | 34 | 31.88 | [9 empty] | 1
| 1 | 2 | 49 | 31.15 | 36.1 | 32.9 | 36.6 | 32.4 | 30.972 | 29.9476 | 39.052 | 31.6 | 25.3 | 48.2785 | 35 | 33.7 | 36.3 | 41.398 | 33 | 33.23 | [9 empty] | 31
| 1 | 3 | 49.5 | 32.85 | 36.2 | 34.3 | 38.5 | 32.9 | 31.237 | 30.2112 | 40.0854 | 33.3 | 26.5 | 48.4153 | 34 | 33 | 34.4 | 41.3 | 34 | 31.9 | [9 empty] | 16
| 1 | 4 | 49.2 | 34 | 36.1 | 32.7 | 40 | 33.7 | 31.393 | 28.4261 | 39.5719 | 33.9 | 25.3 | 48.1245 | 35 | 34.7 | 33.2 | 41.276 | 34 | 32.36 | [9 empty] | 34.0264
| 1 | 5 | 50.6 | 34.06 | 36.9 | 32.7 | 39.5 | 32.2 | 31.517 | 31.7934 | 38.677 | 33.9 | 33.2 | 45.4687 | 34 | 33.2 | 32 | 42.204 | 33 | 33.49 | [9 empty] | 80
| 1 | 6 | 49.4 | 33.39 | 35.8 | 32.5 | 38.4 | 31.1 | 31.976 | 34.5499 | 38.3372 | 33.2 | 26 | 46.417 | 34 | 33.1 | 34.1 | 42.578 | 32 | 33.8
| Mean |  | 49.4667 | 32.8233 | 36.1 | 33.0667 | 38.7833 | 32.6667 | 31.4242 | 30.6313 | 39.1312 | 33.1667 | 27.3833 | 46.9973 | 34.5 | 33.75 | 34.15 | 41.6598 | 33.3333 | 32.7767
| Median |  | 49.3 | 33.12 | 36.1 | 32.8 | 39 | 32.65 | 31.4215 | 30.0794 | 39.0579 | 33.25 | 26.25 | 47.2708 | 34.5 | 33.45 | 34.25 | 41.349 | 33.5 | 32.795
| Std Dev. |  | 0.58538 | 1.2501 | 0.46904 | 0.66232 | 1.2513 | 0.99331 | 0.33295 | 2.2518 | 0.62486 | 0.84301 | 3.0195 | 1.4524 | 0.54772 | 0.81179 | 1.468 | 0.58193 | 0.8165 | 0.83758
| Rel.Std.Dev. |  | 1.183% | 3.809% | 1.299% | 2.003% | 3.226% | 3.041% | 1.06% | 7.351% | 1.597% | 2.542% | 11.03% | 3.09% | 1.588% | 2.405% | 4.299% | 1.397% | 2.449% | 2.555%
| PDM3 |  | 45.38% | -3.536% | 6.094% | -2.821% | 13.98% | -3.996% | -7.648% | -9.978% | 15% | -2.527% | -19.52% | 38.12% | 1.392% | -0.8124% | 0.3632% | 22.43% | -2.037% | -3.673%
| Z-Score (Absolute) |  | 6.98 | 0.31 | 1.12 | 0.21 | 2.3 | 0.38 | 0.93 | 1.28 | 2.45 | 0.16 | 2.7 | 5.9 | 0.42 | 0.09 | 0.27 | 3.56 | 0.09 | 0.34
| Analytical results for Co in OREAS 505b (Certified Value 8.83 ppm) | [28 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 3
| 1 | 1 | 8.5 | 8.8 | 9.2 | 9.6 | 9.3 | 8.9 | 9.7 | 8.4 | 8.91 | 8.0731 | 9.0664 | 10.8 | 8.16 | 11.69 | 11.6 | 8.9 | 8.8634 | 9.1 | 8.2 | 8.8 | 11.8 | 9.1 | 8.622 | 7.7 | 9.2 | 9 |  | 1
| 1 | 2 | 8.4 | 8.5 | 9.2 | 9.2 | 9.2 | 9.2 | 9.5 | 8.3 | 9.1 | 7.6453 | 8.9769 | 10.4 | 8.02 | 12.07 | 11.1 | 8.5 | 8.998 | 8.7 | 8.2 | 9.1 | 12 | 8.5 | 8.484 | 7.7 | 9.5 | 8 |  | 17
| 1 | 3 | 8.2 | 8.7 | 9.3 | 9.2 | 9.3 | 9.4 | 9.7 | 8.4 | 8.85 | 8.0052 | 8.7561 | 10.4 | 8.27 | 11.88 | 11.3 | 9 | 8.7824 | 9.3 | 8.3 | 8.7 | 11.9 | 9.3 | 8.854 | 7.6 | 9 | 9 |  | 16
| 1 | 4 | 8.5 | 8.4 | 9.3 | 9.3 | 8.8 | 9.1 | 10.1 | 8.4 | 8.74 | 8.1808 | 8.9249 | 10.5 | 8.03 | 11.89 | 10.9 | 8.6 | 8.9869 | 9.1 | 8 | 8.8 | 12 | 8.8 | 8.541 | 7.6 | 9.6 | 9 |  | 8.829
| 1 | 5 | 8.4 | 8.6 | 9.2 | 9.1 | 9.1 | 9.1 | 9.5 | 8.4 | 9.06 | 7.4912 | 8.9159 | 10.1 | 8.12 | 12.28 | 10.6 | 9.1 | 9.0779 | 9.2 | 8.8 | 8.5 | 12 | 9.2 | 8.941 | 7.5 | 9.4 | 9 |  | 81
| 1 | 6 | 8.7 | 9.5 | 9.1 | 8.7 | 9 | 8.8 | 9.5 | 8 | 9.04 | 7.7236 | 8.8159 | 10 | 8.11 | 11.92 | 10.8 | 8.8 | 8.8634 | 9.5 | 8.4 | 8.7 | 12.2 | 9.1 | 8.897 | 7.5 | 9.4 | 9
| Mean |  | 8.45 | 8.75 | 9.2167 | 9.1833 | 9.1167 | 9.0833 | 9.6667 | 8.3167 | 8.95 | 7.8532 | 8.9094 | 10.3667 | 8.1183 | 11.955 | 11.05 | 8.8167 | 8.9286 | 9.15 | 8.3167 | 8.7667 | 11.9833 | 9 | 8.7232 | 7.6 | 9.35 | 8.8333
| Median |  | 8.45 | 8.65 | 9.2 | 9.2 | 9.15 | 9.1 | 9.6 | 8.4 | 8.975 | 7.8644 | 8.9204 | 10.4 | 8.115 | 11.905 | 11 | 8.85 | 8.9251 | 9.15 | 8.25 | 8.75 | 12 | 9.1 | 8.738 | 7.6 | 9.4 | 9
| Std Dev. |  | 0.16432 | 0.3937 | 0.075277 | 0.29269 | 0.19408 | 0.2137 | 0.23381 | 0.16021 | 0.14029 | 0.27197 | 0.11109 | 0.28752 | 0.09196 | 0.20007 | 0.36194 | 0.23166 | 0.10989 | 0.26646 | 0.27142 | 0.19664 | 0.13292 | 0.29665 | 0.19769 | 0.089443 | 0.21679 | 0.40825
| Rel.Std.Dev. |  | 1.945% | 4.499% | 0.8168% | 3.187% | 2.129% | 2.353% | 2.419% | 1.926% | 1.567% | 3.463% | 1.247% | 2.773% | 1.133% | 1.674% | 3.275% | 2.628% | 1.231% | 2.912% | 3.264% | 2.243% | 1.109% | 3.296% | 2.266% | 1.177% | 2.319% | 4.622%
| PDM3 |  | -4.292% | -0.8944% | 4.391% | 4.014% | 3.259% | 2.881% | 9.488% | -5.803% | 1.371% | -11.05% | 0.9104% | 17.42% | -8.049% | 35.41% | 25.16% | -0.1393% | 1.129% | 3.636% | -5.803% | -0.7056% | 35.73% | 1.937% | -1.198% | -13.92% | 5.901% | 0.04944%
| Z-Score (Absolute) |  | 1.26 | 0.51 | 0.67 | 0.59 | 0.42 | 0.34 | 1.81 | 1.6 | 0 | 2.77 | 0.1 | 3.58 | 2.1 | 7.6 | 5.31 | 0.34 | 0.05 | 0.51 | 1.6 | 0.46 | 7.67 | 0.13 | 0.57 | 3.41 | 1.01 | NA
Results from laboratory 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Cr in OREAS 505b (Certified Value 55 ppm) | [28 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 0
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  | 1
| 1 | 1 | 52 | 50 | 55.7 | 58 | 55 | 58 | 57 | 55 | 54.8 | 50.6564 | 54.5026 | 56.8 | 53.91 | 55.34 | 56 | 51.0923 | 55 | 54 | 58 | 56.2 | 52.7 | 51.375 | 55 | 57 | 54 |  |  | 1
| 1 | 2 | 51 | 51 | 56.1 | 58 | 55 | 57 | 56 | 55 | 54.9 | 49.8455 | 54.2192 | 55.4 | 51.73 | 55.28 | 57 | 53.1645 | 54 | 54 | 60 | 56.5 | 51.7 | 50.665 | 53 | 57 | 54 |  |  | 18
| 1 | 3 | 51 | 51 | 54.8 | 58 | 54 | 58 | 57 | 55 | 54.5 | 49.7704 | 53.2416 | 57.7 | 53.25 | 56.73 | 58 | 50.6186 | 53 | 54 | 59 | 57.1 | 53.9 | 51.147 | 55 | 58 | 54 |  |  | 16
| 1 | 4 | 51 | 52 | 55.4 | 59 | 55 | 58 | 55 | 56 | 53.4 | 49.8628 | 54.524 | 57.6 | 53.48 | 56.29 | 55 | 53.0072 | 54 | 54 | 56 | 57.1 | 53.7 | 50.028 | 53 | 57 | 53 |  |  | 54.6312
| 1 | 5 | 51 | 52 | 55.6 | 59 | 54 | 57 | 55 | 55 | 55.9 | 50.3906 | 53.7717 | 54.9 | 54.07 | 59.05 | 57 | 53.5724 | 54 | 54 | 56 | 57.1 | 53.3 | 53.981 | 52 | 56 | 54 |  |  | 82
| 1 | 6 | 52 | 51 | 57 | 58 | 54 | 57 | 54 | 56 | 55.5 | 50.0681 | 53.6959 | 57.2 | 56.13 | 57.46 | 55 | 53.7632 | 54 | 53 | 58 | 58.1 | 53.8 | 50.783 | 53 | 59 | 54
| Mean |  | 51.3333 | 51.1667 | 55.7667 | 58.3333 | 54.5 | 57.5 | 55.6667 | 55.3333 | 54.8333 | 50.099 | 53.9925 | 56.6 | 53.7617 | 56.6917 | 56.3333 | 52.5364 | 54 | 53.8333 | 57.8333 | 57.0167 | 53.1833 | 51.3298 | 53.5 | 57.3333 | 53.8333
| Median |  | 51 | 51 | 55.65 | 58 | 54.5 | 57.5 | 55.5 | 55 | 54.85 | 49.9654 | 53.9954 | 57 | 53.695 | 56.51 | 56.5 | 53.0858 | 54 | 54 | 58 | 57.1 | 53.5 | 50.965 | 53 | 57 | 54
| Std Dev. |  | 0.5164 | 0.75277 | 0.73937 | 0.5164 | 0.54772 | 0.54772 | 1.2111 | 0.5164 | 0.86641 | 0.35347 | 0.50885 | 1.1781 | 1.4279 | 1.4237 | 1.2111 | 1.3385 | 0.63246 | 0.40825 | 1.6021 | 0.65243 | 0.84951 | 1.3784 | 1.2247 | 1.0328 | 0.40825
| Rel.Std.Dev. |  | 1.006% | 1.471% | 1.326% | 0.8853% | 1.005% | 0.9526% | 2.176% | 0.9332% | 1.58% | 0.7055% | 0.9424% | 2.082% | 2.656% | 2.511% | 2.15% | 2.548% | 1.171% | 0.7584% | 2.77% | 1.144% | 1.597% | 2.685% | 2.289% | 1.801% | 0.7584%
| PDM3 |  | -6.037% | -6.342% | 2.078% | 6.777% | -0.2402% | 5.251% | 1.895% | 1.285% | 0.3699% | -8.296% | -1.169% | 3.604% | -1.592% | 3.772% | 3.116% | -3.835% | -1.155% | -1.461% | 5.861% | 4.366% | -2.65% | -6.043% | -2.071% | 4.946% | -1.461%
| Z-Score (Absolute) |  | 1.16 | 1.23 | 0.47 | 1.41 | 0 | 1.1 | 0.43 | 0.31 | 0.12 | 1.62 | 0.19 | 0.77 | 0.27 | 0.81 | 0.67 | 0.72 | 0.18 | 0.25 | 1.23 | 0.93 | 0.48 | 1.17 | 0.37 | 1.04 | 0.25
| Analytical results for Cs in OREAS 505b (Certified Value 8.03 ppm) | [28 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.27 | 1.29 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [12 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.25g | [12 empty] | 3
| 1 | 1 | 7.81 | 8.75 | 8.13 | 8.41 | 7.74 | 7.91 | 8.2636 | 10.3 | 10.2 | 7.1489 | 7.6 | 8.08 | 8.15 | 8.2 | 8.3 | [12 empty] | 1
| 1 | 2 | 8.14 | 8.95 | 8.06 | 8.2 | 7.53 | 7.81 | 8.358 | 7.3 | 9.94 | 7.2608 | 7.22 | 8.19 | 8.36 | 8.2 | 8.7 | [12 empty] | 34
| 1 | 3 | 8.22 | 8.87 | 8.31 | 8.27 | 7.61 | 7.76 | 8.0119 | 6.26 | 10.3 | 7.1477 | 7.73 | 7.97 | 8.06 | 8.1 | 8.53 | [12 empty] | 16
| 1 | 4 | 7.91 | 8.83 | 7.95 | 8.54 | 7.64 | 7.85 | 8.1191 | 6.79 | 9.81 | 7.0761 | 7.65 | 7.85 | 7.8 | 8.1 | 8.64 | [12 empty] | 8.0329
| 1 | 5 | 8 | 8.8 | 8.04 | 8.47 | 7.83 | 7.84 | 8.2068 | 7.2 | 10.1 | 7.295 | 7.71 | 7.88 | 7.72 | 8 | 8.5 | [12 empty] | 83
| 1 | 6 | 8.09 | 8.79 | 7.82 | 8.07 | 7.18 | 7.84 | 8.0903 | 6.91 | 9.75 | 7.1687 | 7.9 | 7.89 | 8.14 | 8 | 8.78
| Mean |  | 8.0283 | 8.8317 | 8.0517 | 8.3267 | 7.5883 | 7.835 | 8.175 | 7.46 | 10.0167 | 7.1829 | 7.635 | 7.9767 | 8.0383 | 8.1 | 8.575
| Median |  | 8.045 | 8.815 | 8.05 | 8.34 | 7.625 | 7.84 | 8.1629 | 7.055 | 10.02 | 7.1588 | 7.68 | 7.93 | 8.1 | 8.1 | 8.585
| Std Dev. |  | 0.15198 | 0.070545 | 0.16558 | 0.17784 | 0.22569 | 0.049295 | 0.12589 | 1.4388 | 0.21933 | 0.080777 | 0.2274 | 0.13352 | 0.2387 | 0.089443 | 0.17038
| Rel.Std.Dev. |  | 1.893% | 0.7988% | 2.056% | 2.136% | 2.974% | 0.6292% | 1.54% | 19.29% | 2.19% | 1.125% | 2.978% | 1.674% | 2.969% | 1.104% | 1.987%
| PDM3 |  | -0.05665% | 9.944% | 0.2338% | 3.657% | -5.534% | -2.463% | 1.769% | -7.132% | 24.7% | -10.58% | -4.953% | -0.6998% | 0.06784% | 0.8355% | 6.749%
| Z-Score (Absolute) |  | 0.02 | 1.86 | 0.03 | 0.67 | 1.05 | 0.48 | 0.32 | 1.35 | 4.63 | 2 | 0.94 | 0.14 | 0 | 0.14 | 1.26
| Analytical results for Cu in OREAS 505b (Certified Value 0.322 wt.%) | [28 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 3
| 1 | 1 | 0.302 | 0.29281 | 0.31354 | 0.3255 | 0.342 | 0.328 | 0.333 | 0.317 | 0.32833 | 0.3009 | 0.32535 | 0.326 | 0.33712 | 0.31585 | 0.3237 | 0.3245 | 0.31712 | 0.32311 | 0.323 | 0.341 | 0.3228 | 0.31815 | 0.30642 | 0.32206 | 0.31623 | 0.32683 | 0.3297 | 1
| 1 | 2 | 0.303 | 0.2983 | 0.3098 | 0.3236 | 0.344 | 0.318 | 0.332 | 0.316 | 0.33299 | 0.29869 | 0.31625 | 0.326 | 0.33687 | 0.31371 | 0.3214 | 0.329 | 0.32646 | 0.3158 | 0.322 | 0.344 | 0.3258 | 0.31815 | 0.3154 | 0.30809 | 0.30836 | 0.32892 | 0.3327 | 19
| 1 | 3 | 0.311 | 0.30845 | 0.31254 | 0.3274 | 0.339 | 0.318 | 0.333 | 0.32 | 0.32754 | 0.29779 | 0.31432 | 0.327 | 0.33641 | 0.31097 | 0.3268 | 0.3281 | 0.31198 | 0.32425 | 0.323 | 0.334 | 0.3242 | 0.31714 | 0.30951 | 0.32285 | 0.30722 | 0.32465 | 0.3362 | 16
| 1 | 4 | 0.3 | 0.30376 | 0.30999 | 0.327 | 0.333 | 0.319 | 0.335 | 0.317 | 0.32838 | 0.29833 | 0.32587 | 0.322 | 0.33684 | 0.31129 | 0.3248 | 0.3205 | 0.31557 | 0.31883 | 0.319 | 0.323 | 0.3255 | 0.32118 | 0.30787 | 0.3227 | 0.30782 | 0.33525 | 0.3361 | 0.32236
| 1 | 5 | 0.306 | 0.30746 | 0.30881 | 0.3278 | 0.348 | 0.324 | 0.326 | 0.323 | 0.32638 | 0.30057 | 0.32085 | 0.329 | 0.3301 | 0.32589 | 0.3245 | 0.3283 | 0.32964 | 0.32434 | 0.329 | 0.32 | 0.3246 | 0.32926 | 0.31415 | 0.32201 | 0.3114 | 0.33032 | 0.3244 | 84
| 1 | 6 | 0.31 | 0.29978 | 0.30887 | 0.3243 | 0.331 | 0.317 | 0.327 | 0.326 | 0.3295 | 0.2999 | 0.32135 | 0.324 | 0.3313 | 0.31664 | 0.3277 | 0.3229 | 0.32028 | 0.33108 | 0.317 | 0.325 | 0.327 | 0.33229 | 0.32246 | 0.32453 | 0.3055 | 0.33185 | 0.322
| Mean |  | 0.30533 | 0.30176 | 0.31059 | 0.32593 | 0.3395 | 0.32067 | 0.331 | 0.31983 | 0.32885 | 0.29936 | 0.32066 | 0.32567 | 0.33477 | 0.31572 | 0.32482 | 0.32555 | 0.32017 | 0.3229 | 0.32217 | 0.33117 | 0.32498 | 0.3227 | 0.31263 | 0.32037 | 0.30942 | 0.32964 | 0.33018
| Median |  | 0.3045 | 0.30177 | 0.3099 | 0.32625 | 0.3405 | 0.3185 | 0.3325 | 0.3185 | 0.32835 | 0.2993 | 0.3211 | 0.326 | 0.33662 | 0.31478 | 0.32465 | 0.3263 | 0.3187 | 0.32368 | 0.3225 | 0.3295 | 0.32505 | 0.31966 | 0.31183 | 0.32238 | 0.30809 | 0.32962 | 0.3312
| Std Dev. |  | 0.0044572 | 0.0059538 | 0.001979 | 0.0017363 | 0.0065345 | 0.0043665 | 0.0036332 | 0.0039707 | 0.0022738 | 0.0012729 | 0.0046763 | 0.0024221 | 0.0031864 | 0.0054867 | 0.002248 | 0.0034524 | 0.0067355 | 0.0052503 | 0.0041191 | 0.0099883 | 0.0014511 | 0.006475 | 0.0059603 | 0.0060873 | 0.0038519 | 0.0037452 | 0.0059697
| Rel.Std.Dev. |  | 1.46% | 1.973% | 0.6372% | 0.5327% | 1.925% | 1.362% | 1.098% | 1.241% | 0.6914% | 0.4252% | 1.458% | 0.7437% | 0.9518% | 1.738% | 0.6921% | 1.06% | 2.104% | 1.626% | 1.279% | 3.016% | 0.4465% | 2.007% | 1.906% | 1.9% | 1.245% | 1.136% | 1.808%
| PDM3 |  | -5.281% | -6.389% | -3.649% | 1.11% | 5.318% | -0.5239% | 2.682% | -0.7825% | 2.015% | -7.133% | -0.5253% | 1.027% | 3.852% | -2.057% | 0.7635% | 0.9909% | -0.6768% | 0.1694% | -0.05862% | 2.733% | 0.8152% | 0.1053% | -3.016% | -0.6155% | -4.012% | 2.259% | 2.428%
| Z-Score (Absolute) |  | 1.9 | 2.29 | 1.33 | 0.35 | 1.84 | 0.22 | 0.91 | 0.31 | 0.67 | 2.56 | 0.22 | 0.33 | 1.32 | 0.76 | 0.23 | 0.31 | 0.28 | 0.02 | 0.06 | 0.93 | 0.25 | 0 | 1.1 | 0.25 | 1.45 | 0.76 | 0.82
| Analytical results for Dy in OREAS 505b (Indicative Value 2.51 ppm) | [28 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g | 2
| 1 | 1 | 1.01: 2.376; 1.02: 2.1723; 1.03: 2.9032 | 1
| 1 | 2 | 1.01: 2.399; 1.02: 2.2456; 1.03: 2.8021 | 6
| 1 | 3 | 1.01: 2.455; 1.02: 2.2393; 1.03: 2.7553 | 16
| 1 | 4 | 1.01: 2.364; 1.02: 2.275; 1.03: 2.8125 | 2.5105
| 1 | 5 | 1.01: 2.479; 1.02: 2.4437; 1.03: 2.811 | 12
| 1 | 6 | 1.01: 2.472; 1.02: 2.3558; 1.03: 2.829
| Mean |  | 1.01: 2.4242; 1.02: 2.2886; 1.03: 2.8188
| Median |  | 1.01: 2.427; 1.02: 2.2603; 1.03: 2.8117
| Std Dev. |  | 1.01: 0.050634; 1.02: 0.096501; 1.03: 0.048257
| Rel.Std.Dev. |  | 1.01: 2.089%; 1.02: 4.217%; 1.03: 1.712%
| PDM3 |  | 1.01: -3.44%; 1.02: -8.84%; 1.03: 12.28%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 1.96
| Analytical results for Er in OREAS 505b (Indicative Value 0.97 ppm) | [28 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g | 2
| 1 | 1 | 1.01: 0.9; 1.02: 0.8616; 1.03: 1.1166 | 1
| 1 | 2 | 1.01: 0.909; 1.02: 0.8868; 1.03: 1.1062 | 7
| 1 | 3 | 1.01: 0.908; 1.02: 0.8861; 1.03: 1.0639 | 16
| 1 | 4 | 1.01: 0.902; 1.02: 0.8958; 1.03: 1.0931 | 0.96548
| 1 | 5 | 1.01: 0.933; 1.02: 0.9874; 1.03: 1.0865 | 13
| 1 | 6 | 1.01: 0.933; 1.02: 0.92; 1.03: 1.087
| Mean |  | 1.01: 0.91417; 1.02: 0.90628; 1.03: 1.0922
| Median |  | 1.01: 0.9085; 1.02: 0.8913; 1.03: 1.0901
| Std Dev. |  | 1.01: 0.014986; 1.02: 0.043958; 1.03: 0.01817
| Rel.Std.Dev. |  | 1.01: 1.639%; 1.02: 4.85%; 1.03: 1.664%
| PDM3 |  | 1.01: -5.314%; 1.02: -6.131%; 1.03: 13.13%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 15.23
| Analytical results for Eu in OREAS 505b (Indicative Value 0.43 ppm) | [28 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g | 2
| 1 | 1 | 1.01: 0.359; 1.02: 0.4681; 1.03: 0.45238 | 1
| 1 | 2 | 1.01: 0.362; 1.02: 0.4726; 1.03: 0.44359 | 8
| 1 | 3 | 1.01: 0.35; 1.02: 0.4714; 1.03: 0.44585 | 16
| 1 | 4 | 1.01: 0.346; 1.02: 0.4961; 1.03: 0.43375 | 0.42652
| 1 | 5 | 1.01: 0.36; 1.02: 0.5373; 1.03: 0.42985 | 14
| 1 | 6 | 1.01: 0.365; 1.02: 0.4902; 1.03: 0.45193
| Mean |  | 1.01: 0.357; 1.02: 0.48928; 1.03: 0.44289
| Median |  | 1.01: 0.3595; 1.02: 0.4814; 1.03: 0.44472
| Std Dev. |  | 1.01: 0.0073756; 1.02: 0.026076; 1.03: 0.0093239
| Rel.Std.Dev. |  | 1.01: 2.066%; 1.02: 5.329%; 1.03: 2.105%
| PDM3 |  | 1.01: -16.3%; 1.02: 14.71%; 1.03: 3.838%
| Z-Score (Absolute) |  | 1.01: 1.25; 1.02: 0.67; 1.03: 0
| Analytical results for Fe in OREAS 505b (Certified Value 2.96 wt.%) | [28 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 3
| 1 | 1 | 2.98 | 2.73 | 2.85 | 3.1 | 3.11 | 2.99 | 3.06 | 2.91 | 2.942 | 2.9565 | 2.9292 | 2.83 | 3.82 | 2.9897 | 2.92 | 2.96 | 2.8605 | 2.85 | 2.92 | 3.19 | 3.1 | 2.9 | 3.082 | 2.9138 | 3.03 | 2.94 | 2.85 | 1
| 1 | 2 | 2.98 | 2.77 | 2.83 | 3.1 | 3.07 | 2.99 | 3.04 | 2.93 | 2.952 | 2.9738 | 2.8507 | 2.86 | 3.295 | 2.9795 | 2.95 | 3.01 | 3.0566 | 2.82 | 2.95 | 3.27 | 3.11 | 2.95 | 3.1218 | 2.7858 | 3.01 | 2.99 | 2.87 | 12
| 1 | 3 | 2.94 | 2.8 | 2.83 | 3.11 | 3.07 | 2.98 | 3.05 | 2.94 | 2.916 | 2.9846 | 2.8386 | 2.85 | 3.77 | 2.9243 | 2.95 | 3.05 | 2.9598 | 2.8 | 2.93 | 3.18 | 3.13 | 3.06 | 3.0162 | 2.8734 | 3.06 | 2.97 | 2.81 | 16
| 1 | 4 | 2.96 | 2.82 | 2.82 | 3.13 | 3.08 | 3 | 2.81 | 2.96 | 2.948 | 2.9542 | 2.8747 | 2.82 | 3.99 | 2.9765 | 2.89 | 2.93 | 2.9272 | 2.81 | 2.91 | 3.02 | 3.18 | 3.01 | 3.0358 | 2.9011 | 3.01 | 2.97 | 2.82 | 2.9583
| 1 | 5 | 2.98 | 2.84 | 2.83 | 3.12 | 3.04 | 2.97 | 2.78 | 2.94 | 2.974 | 2.9782 | 2.8457 | 2.9 | 3.98 | 2.9661 | 2.96 | 3.05 | 3.0929 | 2.84 | 2.98 | 3.03 | 3.08 | 3.09 | 3.0817 | 2.8805 | 2.97 | 2.94 | 2.86 | 85
| 1 | 6 | 3.01 | 2.8 | 2.81 | 3.1 | 3.02 | 2.98 | 2.77 | 2.96 | 2.962 | 2.99 | 2.8723 | 2.82 | 3.59 | 2.9587 | 3.02 | 2.95 | 2.9166 | 2.81 | 2.88 | 3.13 | 3.11 | 3.14 | 3.1012 | 2.8143 | 2.99 | 3 | 2.85
| Mean |  | 2.975 | 2.7933 | 2.8283 | 3.11 | 3.065 | 2.985 | 2.9183 | 2.94 | 2.949 | 2.9729 | 2.8685 | 2.8467 | 3.7408 | 2.9658 | 2.9483 | 2.9917 | 2.969 | 2.8217 | 2.9283 | 3.1367 | 3.1183 | 3.025 | 3.0731 | 2.8615 | 3.0117 | 2.9683 | 2.8433
| Median |  | 2.98 | 2.8 | 2.83 | 3.105 | 3.07 | 2.985 | 2.925 | 2.94 | 2.95 | 2.976 | 2.8615 | 2.84 | 3.795 | 2.9713 | 2.95 | 2.985 | 2.9435 | 2.815 | 2.925 | 3.155 | 3.11 | 3.035 | 3.0818 | 2.877 | 3.01 | 2.97 | 2.85
| Std Dev. |  | 0.023452 | 0.038816 | 0.013292 | 0.012649 | 0.031464 | 0.010488 | 0.14497 | 0.018974 | 0.019708 | 0.014657 | 0.033072 | 0.030768 | 0.26379 | 0.023005 | 0.043551 | 0.052313 | 0.088752 | 0.019408 | 0.034303 | 0.097502 | 0.034303 | 0.08961 | 0.039845 | 0.05052 | 0.031252 | 0.024833 | 0.023381
| Rel.Std.Dev. |  | 0.7883% | 1.39% | 0.4699% | 0.4067% | 1.027% | 0.3514% | 4.968% | 0.6454% | 0.6683% | 0.493% | 1.153% | 1.081% | 7.052% | 0.7757% | 1.477% | 1.749% | 2.989% | 0.6878% | 1.171% | 3.108% | 1.1% | 2.962% | 1.297% | 1.766% | 1.038% | 0.8366% | 0.8223%
| PDM3 |  | 0.5659% | -5.575% | -4.392% | 5.129% | 3.608% | 0.9039% | -1.35% | -0.6172% | -0.313% | 0.4944% | -3.033% | -3.772% | 26.45% | 0.2549% | -0.3355% | 1.129% | 0.3614% | -4.617% | -1.012% | 6.031% | 5.411% | 2.256% | 3.882% | -3.271% | 1.805% | 0.3405% | -3.885%
| Z-Score (Absolute) |  | 0.08 | 2.08 | 1.67 | 1.69 | 1.15 | 0.2 | 0.59 | 0.34 | 0.23 | 0.05 | 1.19 | 1.45 | 9.19 | 0.03 | 0.24 | 0.28 | 0.01 | 1.75 | 0.48 | 2 | 1.78 | 0.67 | 1.25 | 1.27 | 0.52 | 0 | 1.49
| Analytical results for Ga in OREAS 505b (Certified Value 8.92 ppm) | [28 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g |  |  |  | 3
| 1 | 1 | 9.6 | 8 | 8.1 | 9.2 | 8.56 | 9.11 | 9.34 | 8.93 | 8.8 | 8.7668 | 9.79 | 10.61 | 11.13 | 9.9 | 11.1 | 8.4851 | 9.36 | 8.18 | 8.87 | 8.3 | 9.3 | 8.42 | 9 | 8.7 |  |  |  | 1
| 1 | 2 | 9.5 | 8 | 8 | 9 | 8.47 | 8.99 | 8.75 | 8.61 | 8.8 | 8.487 | 9.46 | 10.35 | 11.34 | 9.6 | 10.3 | 8.9279 | 9.23 | 8.55 | 9.09 | 8.4 | 9.5 | 8.969 | 9 | 9.1 |  |  |  | 39
| 1 | 3 | 9.3 | 8 | 8.1 | 9.2 | 8.83 | 9.2 | 9.49 | 8.82 | 8.9 | 8.4831 | 9.55 | 10.56 | 11.53 | 9.7 | 10.8 | 8.4428 | 9.33 | 8.1 | 8.81 | 8.4 | 9.8 | 9.003 | 9 | 8.9 |  |  |  | 16
| 1 | 4 | 9.2 | 8 | 7.9 | 9.1 | 8.48 | 8.98 | 9.27 | 8.78 | 8.9 | 8.6452 | 9.64 | 10.51 | 11.46 | 9.6 | 10.3 | 8.549 | 9.21 | 8.34 | 8.57 | 8.5 | 9.2 | 9.164 | 9 | 8.9 |  |  |  | 8.9227
| 1 | 5 | 9.5 | 8 | 8 | 8.8 | 8.75 | 9.03 | 8.93 | 8.81 | 9.2 | 8.4299 | 9.6 | 10.59 | 12.22 | 9.3 | 10.4 | 8.8779 | 9.34 | 8.4 | 8.46 | 8.5 | 9.8 | 9.373 | 9 | 9.1 |  |  |  | 86
| 1 | 6 | 9.5 | 9 | 8.1 | 8.9 | 8.61 | 8.73 | 8.85 | 8.12 | 9 | 8.4805 | 9.46 | 10.75 | 11.76 | 9.5 | 10.6 | 8.5269 | 9.45 | 8.38 | 8.88 | 8.5 | 9.5 | 8.76 | 9 | 8.9
| Mean |  | 9.4333 | 8.1667 | 8.0333 | 9.0333 | 8.6167 | 9.0067 | 9.105 | 8.6783 | 8.9333 | 8.5488 | 9.5833 | 10.5617 | 11.5733 | 9.6 | 10.5833 | 8.6349 | 9.32 | 8.325 | 8.78 | 8.4333 | 9.5167 | 8.9482 | 9 | 8.9333
| Median |  | 9.5 | 8 | 8.05 | 9.05 | 8.585 | 9.01 | 9.1 | 8.795 | 8.9 | 8.4851 | 9.575 | 10.575 | 11.495 | 9.6 | 10.5 | 8.5379 | 9.335 | 8.36 | 8.84 | 8.45 | 9.5 | 8.986 | 9 | 8.9
| Std Dev. |  | 0.15055 | 0.40825 | 0.08165 | 0.1633 | 0.1461 | 0.15883 | 0.30078 | 0.2924 | 0.15055 | 0.12944 | 0.12469 | 0.13121 | 0.37914 | 0.2 | 0.31885 | 0.21133 | 0.088544 | 0.16196 | 0.22874 | 0.08165 | 0.24833 | 0.33011 | 0 | 0.15055
| Rel.Std.Dev. |  | 1.596% | 4.999% | 1.016% | 1.808% | 1.696% | 1.763% | 3.303% | 3.369% | 1.685% | 1.514% | 1.301% | 1.242% | 3.276% | 2.083% | 3.013% | 2.447% | 0.95% | 1.945% | 2.605% | 0.9682% | 2.609% | 3.689% | 0% | 1.685%
| PDM3 |  | 5.723% | -8.473% | -9.968% | 1.24% | -3.43% | 0.9409% | 2.043% | -2.739% | 0.119% | -4.191% | 7.404% | 18.37% | 29.71% | 7.591% | 18.61% | -3.225% | 4.453% | -6.699% | -1.599% | -5.485% | 6.657% | 0.2853% | 0.8662% | 0.119%
| Z-Score (Absolute) |  | 0.78 | NA | 1.61 | 0.1 | 0.62 | 0.05 | 0.22 | 0.51 | 0.08 | 0.73 | 1.04 | 2.71 | 4.43 | 1.06 | 2.74 | 0.59 | 0.59 | 1.11 | 0.34 | 0.93 | 0.92 | 0.05 | NA | 0.08
Results from laboratories 1.02 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 505b (Indicative Value 3.46 ppm) | [28 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g |  |  |  | 2
| 1 | 1 | 1.01: 3.318; 1.02: 3.1375; 1.03: 3.7527 |  |  |  | 1
| 1 | 2 | 1.01: 3.328; 1.02: 3.3296; 1.03: 3.5998 |  |  |  | 9
| 1 | 3 | 1.01: 3.408; 1.02: 3.2423; 1.03: 3.5683 |  |  |  | 16
| 1 | 4 | 1.01: 3.295; 1.02: 3.3622; 1.03: 3.6201 |  |  |  | 3.4573
| 1 | 5 | 1.01: 3.417; 1.02: 3.6786; 1.03: 3.4912 |  |  |  | 15
| 1 | 6 | 1.01: 3.448; 1.02: 3.5829; 1.03: 3.6516
| Mean |  | 1.01: 3.369; 1.02: 3.3889; 1.03: 3.6139
| Median |  | 1.01: 3.368; 1.02: 3.3459; 1.03: 3.6099
| Std Dev. |  | 1.01: 0.062967; 1.02: 0.20518; 1.03: 0.08725
| Rel.Std.Dev. |  | 1.01: 1.869%; 1.02: 6.054%; 1.03: 2.414%
| PDM3 |  | 1.01: -2.553%; 1.02: -1.979%; 1.03: 4.532%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0; 1.03: 7.65
| Analytical results for Ge in OREAS 505b (Certified Value 0.1 ppm) | [28 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab |  |  |  | 1
| No | No | 1.01: 1.05; 1.02: 1.06; 1.03: 1.07; 1.04: 1.08; 1.05: 1.09; 1.06: 1.11; 1.07: 1.12; 1.08: 1.17; 1.09: 1.19; 1.11: 1.20; 1.12: 1.21; 1.13: 1.27; 1.14: 1.29 |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS; 1.09: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.2g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 0.2g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.25g; 1.14: 0.25g |  |  |  | 2
| 1 | 1 | 1.01: 0.11; 1.02: 0.09; 1.03: 0.13; 1.04: 0.14; 1.05: 0.09; 1.06: < 0.2; 1.07: < 0.05; 1.08: < 0.1; 1.09: 0.12; 1.11: 0.12; 1.12: 0.1; 1.13: 0.9; 1.14: 0.2 |  |  |  | 1
| 1 | 2 | 1.01: 0.1; 1.02: 0.09; 1.03: 0.13; 1.04: 0.13; 1.05: 0.08; 1.06: < 0.2; 1.07: 0.4; 1.08: < 0.1; 1.09: 0.1; 1.11: 0.11; 1.12: 0.09; 1.13: 0.9; 1.14: 0.2 |  |  |  | 4
| 1 | 3 | 1.01: 0.1; 1.02: 0.09; 1.03: 0.14; 1.04: 0.14; 1.05: 0.08; 1.06: < 0.2; 1.07: 0.3; 1.08: < 0.1; 1.09: 0.12; 1.11: 0.12; 1.12: 0.09; 1.13: 0.9; 1.14: 0.2 |  |  |  | 16
| 1 | 4 | 1.01: 0.1; 1.02: 0.09; 1.03: 0.09; 1.04: 0.11; 1.05: 0.08; 1.06: < 0.2; 1.07: 0.66; 1.08: < 0.1; 1.09: 0.12; 1.11: 0.12; 1.12: 0.07; 1.13: 0.9; 1.14: 0.2 |  |  |  | 0.10417
| 1 | 5 | 1.01: 0.1; 1.02: 0.1; 1.03: 0.08; 1.04: 0.13; 1.05: 0.08; 1.06: < 0.2; 1.07: 0.27; 1.08: < 0.1; 1.09: 0.12; 1.11: 0.13; 1.12: 0.09; 1.13: 0.8; 1.14: 0.2 |  |  |  | 87
| 1 | 6 | 1.01: 0.1; 1.02: 0.09; 1.03: 0.08; 1.04: 0.12; 1.05: 0.08; 1.06: < 0.2; 1.07: 0.54; 1.08: < 0.1; 1.09: 0.13; 1.11: 0.12; 1.12: 0.06; 1.13: 0.9; 1.14: 0.2
| Mean |  | 1.01: 0.10167; 1.02: 0.091667; 1.03: 0.10833; 1.04: 0.12833; 1.05: 0.081667; 1.07: 0.434; 1.09: 0.11833; 1.11: 0.12; 1.12: 0.083333; 1.13: 0.88333; 1.14: 0.2
| Median |  | 1.01: 0.1; 1.02: 0.09; 1.03: 0.11; 1.04: 0.13; 1.05: 0.08; 1.07: 0.4; 1.09: 0.12; 1.11: 0.12; 1.12: 0.09; 1.13: 0.9; 1.14: 0.2
| Std Dev. |  | 1.01: 0.0040825; 1.02: 0.0040825; 1.03: 0.027869; 1.04: 0.01169; 1.05: 0.0040825; 1.07: 0.16456; 1.09: 0.0098319; 1.11: 0.0063246; 1.12: 0.015055; 1.13: 0.040825; 1.14: 3.0405e-17
| Rel.Std.Dev. |  | 1.01: 4.016%; 1.02: 4.454%; 1.03: 25.72%; 1.04: 9.109%; 1.05: 4.999%; 1.07: 37.92%; 1.09: 8.309%; 1.11: 5.27%; 1.12: 18.07%; 1.13: 4.622%; 1.14: 1.52e-14%
| PDM3 |  | 1.01: -2.4%; 1.02: -12%; 1.03: 4%; 1.04: 23.2%; 1.05: -21.6%; 1.07: 316.6%; 1.09: 13.6%; 1.11: 15.2%; 1.12: -20%; 1.13: 748%; 1.14: 92%
| Z-Score (Absolute) |  | 1.01: 0.12; 1.02: 0.49; 1.03: 0.12; 1.04: 0.86; 1.05: 0.86; 1.06: 0.18; 1.07: 9.59; 1.08: 2.02; 1.09: 0.49; 1.11: 0.55; 1.12: 0.8; 1.13: NA; 1.14: 3.49
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 505b (Certified Value 0.32 ppm) | [28 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.27 | 1.29 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [12 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.25g | 0.25g | [12 empty] | 3
| 1 | 1 | 0.37 | 0.32 | 0.3 | 0.31 | 0.31 | 0.31 | 0.40348 | 0.81 | 0.32 | 0.34157 | 0.3 | 0.29 | 0.33 | 0.4 | 0.3 | [12 empty] | 1
| 1 | 2 | 0.36 | 0.3 | 0.31 | 0.3 | 0.31 | 0.29 | 0.42162 | 0.46 | 0.32 | 0.34442 | 0.3 | 0.28 | 0.34 | 0.4 | 0.31 | [12 empty] | 25
| 1 | 3 | 0.34 | 0.3 | 0.3 | 0.32 | 0.33 | 0.3 | 0.3381 | 0.42 | 0.33 | 0.36169 | 0.31 | 0.27 | 0.36 | 0.4 | 0.3 | [12 empty] | 16
| 1 | 4 | 0.33 | 0.29 | 0.3 | 0.32 | 0.32 | 0.3 | 0.3975 | 0.41 | 0.31 | 0.35299 | 0.31 | 0.3 | 0.33 | 0.4 | 0.3 | [12 empty] | 0.31635
| 1 | 5 | 0.33 | 0.3 | 0.3 | 0.33 | 0.34 | 0.29 | 0.38753 | 0.4 | 0.33 | 0.3628 | 0.32 | 0.29 | 0.33 | 0.4 | 0.32 | [12 empty] | 88
| 1 | 6 | 0.34 | 0.31 | 0.31 | 0.33 | 0.3 | 0.28 | 0.3804 | 0.38 | 0.33 | 0.34363 | 0.3 | 0.28 | 0.35 | 0.4 | 0.31
| Mean |  | 0.345 | 0.30333 | 0.30333 | 0.31833 | 0.31833 | 0.295 | 0.3881 | 0.48 | 0.32333 | 0.35118 | 0.30667 | 0.285 | 0.34 | 0.4 | 0.30667
| Median |  | 0.34 | 0.3 | 0.3 | 0.32 | 0.315 | 0.295 | 0.39251 | 0.415 | 0.325 | 0.34871 | 0.305 | 0.285 | 0.335 | 0.4 | 0.305
| Std Dev. |  | 0.016432 | 0.010328 | 0.005164 | 0.01169 | 0.01472 | 0.010488 | 0.028313 | 0.16383 | 0.008165 | 0.0094223 | 0.008165 | 0.010488 | 0.012649 | 6.0809e-17 | 0.008165
| Rel.Std.Dev. |  | 4.763% | 3.405% | 1.702% | 3.672% | 4.624% | 3.555% | 7.295% | 34.13% | 2.525% | 2.683% | 2.662% | 3.68% | 3.72% | 1.52e-14% | 2.662%
| PDM3 |  | 9.057% | -4.114% | -4.114% | 0.6274% | 0.6274% | -6.748% | 22.68% | 51.73% | 2.208% | 11.01% | -3.061% | -9.91% | 7.476% | 26.44% | -3.061%
| Z-Score (Absolute) |  | 0.98 | 0.55 | 0.55 | 0 | 0 | 0.86 | 2.57 | 5.95 | 0.18 | 1.21 | 0.43 | 1.23 | 0.8 | NA | 0.43
Results from laboratory 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hg in OREAS 505b (Certified Value 0.021 ppm) | [28 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.29 | 1.30 | [5 empty] | ppm
|  |  | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | [5 empty] | 3
|  |  | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | [5 empty] | 3
| 1 | 1 | 0.02 | 0.017 | 0.03 | 0.02 | 0.02 | 0.03 | 0.02 | 0.02 | < 1 | 0.49 | < 5 | < 0.1 | < 0.05 | < 2 | 0.02 | 0.04 | < 5 | 0.02 | 0.062 | 0.384 | < 0.01 | 2 | [5 empty] | 1
| 1 | 2 | 0.01 | 0.017 | 0.03 | 0.03 | 0.03 | 0.02 | 0.01 | 0.02 | < 1 | 0.06 | < 5 | < 0.1 | < 0.05 | < 2 | 0.02 | 0.05 | < 5 | 0.033 | 0.056 | 0.399 | < 0.01 | 2 | [5 empty] | 3
| 1 | 3 | 0.02 | 0.017 | 0.02 | 0.02 | 0.03 | 0.02 | 0.03 | 0.02 | < 1 | 0.05 | < 5 | < 0.1 | < 0.05 | < 2 | 0.02 | 0.04 | < 5 | < 0.005 | 0.056 | 0.412 | < 0.01 | 2 | [5 empty] | 16
| 1 | 4 | 0.02 | 0.017 | 0.02 | 0.03 | 0.03 | 0.01 | 0.04 | 0.02 | < 1 | 0.02 | < 5 | < 0.1 | < 0.05 | < 2 | 0.02 | 0.04 | < 5 | 0.009 | 0.055 | 0.395 | < 0.01 | 2 | [5 empty] | 0.020987
| 1 | 5 | 0.03 | 0.014 | 0.03 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 | < 1 | 0.05 | < 5 | < 0.1 | < 0.05 | < 2 | 0.02 | 0.04 | < 5 | 0.009 | 0.058 | 0.417 | < 0.01 | 2 | [5 empty] | 89
| 1 | 6 | 0.01 | 0.015 | 0.03 | 0.02 | 0.03 | 0.01 | 0.01 | 0.02 | < 1 | 0.04 | < 5 | < 0.1 | < 0.05 | < 2 | 0.02 | 0.04 | < 5 | 0.012 | 0.059 | 0.396 | < 0.01 | 2
| Mean |  | 0.018333 | 0.016167 | 0.026667 | 0.023333 | 0.028333 | 0.018333 | 0.021667 | 0.02 |  | 0.11833 |  |  |  |  | 0.02 | 0.041667 |  | 0.0166 | 0.057667 | 0.4005 |  | 2
| Median |  | 0.02 | 0.017 | 0.03 | 0.02 | 0.03 | 0.02 | 0.02 | 0.02 |  | 0.05 |  |  |  |  | 0.02 | 0.04 |  | 0.012 | 0.057 | 0.3975 |  | 2
| Std Dev. |  | 0.0075277 | 0.0013292 | 0.005164 | 0.005164 | 0.0040825 | 0.0075277 | 0.01169 | 0 |  | 0.18258 |  |  |  |  | 0 | 0.0040825 |  | 0.010213 | 0.002582 | 0.012079 |  | 0
| Rel.Std.Dev. |  | 41.06% | 8.222% | 19.36% | 22.13% | 14.41% | 41.06% | 53.96% | 0% |  | 154.3% |  |  |  |  | 0% | 9.798% |  | 61.52% | 4.477% | 3.016% |  | 0%
| PDM3 |  | -12.64% | -22.97% | 27.06% | 11.18% | 35.01% | -12.64% | 3.24% | -4.701% |  | 463.9% |  |  |  |  | -4.701% | 98.54% |  | -20.9% | 174.8% | 1808% |  | 9430%
| Z-Score (Absolute) |  | 0.45 | 0.57 | 0 | 0.18 | 0.09 | 0.45 | 0.27 | 0.36 | 25.71 | 4.98 | 134.32 | 1.27 | 0.09 | 52.86 | 0.36 | 0.81 | 134.32 | 0.67 | 1.68 | 20.3 | 1.18 | NA
Results from laboratory 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 505b (Indicative Value 0.37 ppm) | [28 empty] | Indicative
Ho | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab | [5 empty] | 1
| No | No | 1.02: 1.09; 1.03: 1.10; 1.04: 1.11 | [5 empty] | ppm
|  |  | 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | [5 empty] | 2
|  |  | 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.2g | [5 empty] | 2
| 1 | 1 | 1.02: 0.365; 1.03: 0.3078; 1.04: 0.41127 | [5 empty] | 1
| 1 | 2 | 1.02: 0.371; 1.03: 0.3167; 1.04: 0.42131 | [5 empty] | 10
| 1 | 3 | 1.02: 0.379; 1.03: 0.3164; 1.04: 0.40934 | [5 empty] | 16
| 1 | 4 | 1.02: 0.364; 1.03: 0.3314; 1.04: 0.41045 | [5 empty] | 0.37017
| 1 | 5 | 1.02: 0.382; 1.03: 0.357; 1.04: 0.40287 | [5 empty] | 16
| 1 | 6 | 1.02: 0.378; 1.03: 0.3346; 1.04: 0.40483
| Mean |  | 1.02: 0.37317; 1.03: 0.32732; 1.04: 0.41001
| Median |  | 1.02: 0.3745; 1.03: 0.32405; 1.04: 0.4099
| Std Dev. |  | 1.02: 0.0076267; 1.03: 0.017684; 1.04: 0.006445
| Rel.Std.Dev. |  | 1.02: 2.044%; 1.03: 5.403%; 1.04: 1.572%
| PDM3 |  | 1.02: 0.8109%; 1.03: -11.58%; 1.04: 10.76%
| Z-Score (Absolute) |  | 1.02: 0; 1.03: 0.84; 1.04: 0.67
| Analytical results for In in OREAS 505b (Certified Value 0.077 ppm) | [28 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.12 | 1.13 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.29 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | [11 empty] | 3
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | [11 empty] | 3
| 1 | 1 | 0.08 | 0.072 | 0.076 | 0.075 | 0.076 | 0.07 | 1.69 | < 0.1 | 0.08 | 0.080478 | 0.067 | 0.08 | 0.083 | < 5 | 0.09 | 0.08 | [11 empty] | 1
| 1 | 2 | 0.08 | 0.076 | 0.075 | 0.07 | 0.078 | 0.07 | 0.177 | < 0.1 | 0.08 | 0.079342 | 0.067 | 0.077 | 0.081 | < 5 | 0.09 | 0.08 | [11 empty] | 27
| 1 | 3 | 0.08 | 0.076 | 0.075 | 0.074 | 0.08 | 0.07 | 0.075 | < 0.1 | 0.08 | 0.071726 | 0.075 | 0.077 | 0.081 | < 5 | 0.09 | 0.07 | [11 empty] | 16
| 1 | 4 | 0.08 | 0.07 | 0.078 | 0.077 | 0.077 | 0.07 | 0.072 | < 0.1 | 0.08 | 0.078224 | 0.071 | 0.08 | 0.074 | < 5 | 0.09 | 0.08 | [11 empty] | 0.076657
| 1 | 5 | 0.08 | 0.077 | 0.079 | 0.074 | 0.083 | 0.07 | 0.01 | < 0.1 | 0.08 | 0.073233 | 0.067 | 0.077 | 0.079 | < 5 | 0.09 | 0.07 | [11 empty] | 90
| 1 | 6 | 0.08 | 0.081 | 0.078 | 0.079 | 0.071 | 0.07 | < 0.005 | < 0.1 | 0.08 | 0.083267 | 0.067 | 0.078 | 0.078 | < 5 | 0.09 | 0.08
| Mean |  | 0.08 | 0.075333 | 0.076833 | 0.074833 | 0.0775 | 0.07 | 0.4048 |  | 0.08 | 0.077712 | 0.069 | 0.078167 | 0.079333 |  | 0.09 | 0.076667
| Median |  | 0.08 | 0.076 | 0.077 | 0.0745 | 0.0775 | 0.07 | 0.075 |  | 0.08 | 0.078783 | 0.067 | 0.0775 | 0.08 |  | 0.09 | 0.08
| Std Dev. |  | 0 | 0.0038816 | 0.0017224 | 0.0030605 | 0.0040373 | 0 | 0.72094 |  | 0 | 0.0044118 | 0.0033466 | 0.001472 | 0.0031411 |  | 1.5202e-17 | 0.005164
| Rel.Std.Dev. |  | 0% | 5.153% | 2.242% | 4.09% | 5.209% | 0% | 178.1% |  | 0% | 5.677% | 4.85% | 1.883% | 3.959% |  | 1.689e-14% | 6.736%
| PDM3 |  | 4.362% | -1.726% | 0.2306% | -2.378% | 1.1% | -8.684% | 428.1% |  | 4.362% | 1.377% | -9.988% | 1.97% | 3.492% |  | 17.41% | 0.01321%
| Z-Score (Absolute) |  | 0.67 | 0.64 | 0.22 | 0.78 | 0.03 | 2.14 | 73.27 | 7.78 | 0.67 | 0.03 | 2.42 | 0.16 | 0.49 | 682.28 | 3.49 | 0.26
| Analytical results for K in OREAS 505b (Certified Value 0.911 wt.%) | [28 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 3
| 1 | 1 | 0.915 | 0.8 | 0.93 | 0.93 | 0.96 | 0.98 | 0.95 | 0.9 | 0.898 | 0.90231 | 0.922 | 0.86 | 1.1249 | 0.83858 | 0.9 | 1.01 | 0.87843 | 0.87 | 0.89 | 0.99 | 0.97 | 0.82 | 0.85953 | 0.9 | 0.99 | 0.89 |  | 1
| 1 | 2 | 0.915 | 0.82 | 0.92 | 0.93 | 0.94 | 0.96 | 0.96 | 0.91 | 0.887 | 0.9022 | 0.9018 | 0.86 | 0.96329 | 0.83769 | 0.9 | 1.02 | 0.88194 | 0.88 | 0.89 | 1.01 | 0.98 | 0.82 | 0.86035 | 0.9 | 1 | 0.9 |  | #N/A
| 1 | 3 | 0.926 | 0.84 | 0.92 | 0.93 | 0.94 | 0.98 | 0.95 | 0.9 | 0.88 | 0.91291 | 0.90216 | 0.85 | 1.1337 | 0.8311 | 0.91 | 1.04 | 0.89271 | 0.87 | 0.88 | 0.98 | 0.98 | 0.85 | 0.86729 | 0.9 | 0.99 | 0.88 |  | 16
| 1 | 4 | 0.933 | 0.84 | 0.92 | 0.94 | 0.94 | 0.96 | 0.94 | 0.91 | 0.863 | 0.90459 | 0.9156 | 0.85 | 1.1835 | 0.8304 | 0.9 | 1.01 | 0.88458 | 0.87 | 0.88 | 0.94 | 0.97 | 0.85 | 0.86877 | 0.89 | 1.02 | 0.88 |  | 0.91056
| 1 | 5 | 0.911 | 0.85 | 0.92 | 0.94 | 0.94 | 0.96 | 0.93 | 0.91 | 0.892 | 0.91163 | 0.90518 | 0.86 | 1.174 | 0.83351 | 0.89 | 1.04 | 0.89552 | 0.88 | 0.9 | 0.93 | 0.98 | 0.86 | 0.89199 | 0.86 | 0.99 | 0.9 |  | 91
| 1 | 6 | 0.94 | 0.83 | 0.92 | 0.94 | 0.93 | 0.96 | 0.92 | 0.92 | 0.903 | 0.90525 | 0.90639 | 0.85 | 1.0599 | 0.83631 | 0.9 | 1.01 | 0.89086 | 0.87 | 0.87 | 0.97 | 0.98 | 0.88 | 0.88988 | 0.87 | 0.96 | 0.9
| Mean |  | 0.92333 | 0.83 | 0.92167 | 0.935 | 0.94167 | 0.96667 | 0.94167 | 0.90833 | 0.88717 | 0.90648 | 0.90885 | 0.855 | 1.1065 | 0.8346 | 0.9 | 1.0217 | 0.88734 | 0.87333 | 0.885 | 0.97 | 0.97667 | 0.84667 | 0.87297 | 0.88667 | 0.99167 | 0.89167
| Median |  | 0.9205 | 0.835 | 0.92 | 0.935 | 0.94 | 0.96 | 0.945 | 0.91 | 0.8895 | 0.90492 | 0.90578 | 0.855 | 1.1293 | 0.83491 | 0.9 | 1.015 | 0.88772 | 0.87 | 0.885 | 0.975 | 0.98 | 0.85 | 0.86803 | 0.895 | 0.99 | 0.895
| Std Dev. |  | 0.01157 | 0.017889 | 0.0040825 | 0.0054772 | 0.0098319 | 0.010328 | 0.01472 | 0.0075277 | 0.014331 | 0.0046612 | 0.0081518 | 0.0054772 | 0.082768 | 0.0034465 | 0.0063246 | 0.01472 | 0.0066986 | 0.005164 | 0.010488 | 0.030332 | 0.005164 | 0.023381 | 0.014404 | 0.017512 | 0.019408 | 0.0098319
| Rel.Std.Dev. |  | 1.253% | 2.155% | 0.4429% | 0.5858% | 1.044% | 1.068% | 1.563% | 0.8287% | 1.615% | 0.5142% | 0.8969% | 0.6406% | 7.48% | 0.4129% | 0.7027% | 1.441% | 0.7549% | 0.5913% | 1.185% | 3.127% | 0.5287% | 2.762% | 1.65% | 1.975% | 1.957% | 1.103%
| PDM3 |  | 1.402% | -8.848% | 1.219% | 2.684% | 3.416% | 6.161% | 3.416% | -0.2449% | -2.569% | -0.4483% | -0.1877% | -6.102% | 21.52% | -8.343% | -1.16% | 12.2% | -2.55% | -4.089% | -2.807% | 6.527% | 7.26% | -7.017% | -4.129% | -2.624% | 8.907% | -2.075%
| Z-Score (Absolute) |  | 0.31 | 1.53 | 0.28 | 0.54 | 0.68 | 1.17 | 0.68 | 0.02 | 0.4 | 0.02 | 0.03 | 1.03 | 3.93 | 1.44 | 0.15 | 2.26 | 0.4 | 0.67 | 0.44 | 1.24 | 1.37 | 1.2 | 0.68 | 0.41 | 1.66 | 0.31
| Analytical results for La in OREAS 505b (Certified Value 16.6 ppm) | [28 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  | 1
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  |  | 2
| 1 | 1 | 23.8 | 10 | 14 | 15.8 | 17.4 | 16.2 | 18.5 | 17 | 15.524 | 13.6624 | 16.9765 | 16.5 | 20.7 | 16.3 | 23.1315 | 18.3 | 17.3 | 17.1 | 17.1 | 11.6 | 21.042 | 15 | 14.9 | 17.6 |  |  |  | 1
| 1 | 2 | 23.7 | 10 | 13.5 | 15.9 | 17.5 | 16.2 | 18.2 | 16.3 | 15.362 | 14.023 | 16.9635 | 15.7 | 21.1 | 16.6 | 23.9999 | 17.4 | 17.4 | 17.8 | 17.2 | 12 | 21.099 | 15 | 15.8 | 17.6 |  |  |  | 28
| 1 | 3 | 24 | 14 | 13.9 | 16.5 | 18.2 | 17 | 19 | 16.6 | 15.306 | 14.1797 | 17.4337 | 16.3 | 21.6 | 17.4 | 24.0431 | 18.3 | 16.3 | 17 | 17.3 | 12.4 | 21.277 | 15 | 15.2 | 16.9 |  |  |  | 16
| 1 | 4 | 23.5 | 11 | 14.2 | 17.4 | 17.2 | 16.2 | 19.1 | 17.2 | 15.058 | 13.4673 | 17.384 | 16.9 | 20.3 | 16.6 | 24.129 | 18.2 | 17 | 16.2 | 17.3 | 11.7 | 20.976 | 15 | 15.4 | 17 |  |  |  | 16.5502
| 1 | 5 | 22.5 | 14 | 14 | 17 | 18.6 | 16 | 18.6 | 16.4 | 15.634 | 15.0284 | 16.9634 | 16.5 | 20 | 17.4 | 24.702 | 17.2 | 16.6 | 15.8 | 17.1 | 12.5 | 21.617 | 15 | 15.7 | 17.1 |  |  |  | 92
| 1 | 6 | 24.4 | 10 | 13.7 | 16.7 | 18.2 | 16 | 18.7 | 15.8 | 15.733 | 16.2308 | 16.8746 | 16.3 | 21.3 | 17 | 23.0773 | 18.3 | 16.6 | 16.9 | 17.4 | 11.8 | 21.759 | 15 | 15.6 | 17.1
| Mean |  | 23.65 | 11.5 | 13.8833 | 16.55 | 17.85 | 16.2667 | 18.6833 | 16.55 | 15.4362 | 14.4319 | 17.0993 | 16.3667 | 20.8333 | 16.8833 | 23.8471 | 17.95 | 16.8667 | 16.8 | 17.2333 | 12 | 21.295 | 15 | 15.4333 | 17.2167
| Median |  | 23.75 | 10.5 | 13.95 | 16.6 | 17.85 | 16.2 | 18.65 | 16.5 | 15.443 | 14.1014 | 16.97 | 16.4 | 20.9 | 16.8 | 24.0215 | 18.25 | 16.8 | 16.95 | 17.25 | 11.9 | 21.188 | 15 | 15.5 | 17.1
| Std Dev. |  | 0.64109 | 1.9748 | 0.24833 | 0.62209 | 0.55767 | 0.37238 | 0.33116 | 0.50498 | 0.24512 | 1.0339 | 0.24305 | 0.39328 | 0.6121 | 0.45789 | 0.62878 | 0.50892 | 0.43665 | 0.70711 | 0.12111 | 0.37417 | 0.32357 | 0 | 0.33862 | 0.30605
| Rel.Std.Dev. |  | 2.711% | 17.17% | 1.789% | 3.759% | 3.124% | 2.289% | 1.772% | 3.051% | 1.588% | 7.164% | 1.421% | 2.403% | 2.938% | 2.712% | 2.637% | 2.835% | 2.589% | 4.209% | 0.7027% | 3.118% | 1.519% | 0% | 2.194% | 1.778%
| PDM3 |  | 42.9% | -30.51% | -16.11% | -0.001433% | 7.853% | -1.713% | 12.89% | -0.001433% | -6.731% | -12.8% | 3.317% | -1.109% | 25.88% | 2.013% | 44.09% | 8.458% | 1.912% | 1.509% | 4.127% | -27.49% | 28.67% | -9.367% | -6.749% | 4.027%
| Z-Score (Absolute) |  | 4.46 | NA | 1.97 | 0.21 | 0.64 | 0.4 | 1.19 | 0.21 | 0.95 | 1.61 | 0.15 | 0.33 | 2.6 | 0.01 | 4.59 | 0.71 | 0.01 | 0.05 | 0.24 | 3.21 | 2.91 | NA | 0.95 | 0.22
Results from laboratories 1.02 and 1.27 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 505b (Certified Value 37.6 ppm) | [28 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | [6 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | [6 empty] | 2
| 1 | 1 | 37.8 | 38 | 39.4 | 38.5 | 39.5 | 36.4 | 35.14 | 35.4636 | 35.5783 | 36 | 50.46 | 40 | 34.4137 | 36 | 37.1 | 41.9 | 42.4 | 39.275 | 35.1 | 40 | 36 | [6 empty] | 1
| 1 | 2 | 37.3 | 37 | 38.8 | 38.2 | 39.3 | 37.1 | 34.75 | 36.2776 | 34.9096 | 34.4 | 50.31 | 39 | 37.028 | 36 | 37.7 | 42.8 | 42.6 | 39.024 | 33.2 | 40 | 37 | [6 empty] | #N/A
| 1 | 3 | 36.3 | 38 | 38.7 | 38.2 | 39.4 | 36.7 | 34.95 | 35.8466 | 35.7957 | 36.6 | 49.97 | 41 | 35.3677 | 36 | 37.2 | 41.6 | 42.5 | 39.313 | 33.7 | 39 | 36 | [6 empty] | 16
| 1 | 4 | 36.6 | 37 | 38.9 | 37.9 | 39 | 37 | 34.97 | 35.2385 | 36.171 | 35.5 | 50.21 | 38 | 35.669 | 36 | 37.2 | 39.6 | 42.7 | 39.003 | 33.7 | 40 | 36 | [6 empty] | 37.5627
| 1 | 5 | 36.7 | 38 | 38.8 | 37.7 | 38.3 | 37.3 | 34.99 | 36.5178 | 34.4866 | 35.2 | 49.61 | 40 | 36.7902 | 37 | 38 | 39.8 | 42.3 | 39.999 | 33.3 | 40 | 37 | [6 empty] | 93
| 1 | 6 | 36.8 | 38 | 38.3 | 37.7 | 38.3 | 37.4 | 35.49 | 36.1944 | 34.987 | 36.1 | 49.42 | 38 | 35.1036 | 36 | 36.8 | 41 | 42.2 | 39.782 | 34.5 | 41 | 37
| Mean |  | 36.9167 | 37.6667 | 38.8167 | 38.0333 | 38.9667 | 36.9833 | 35.0483 | 35.9231 | 35.3214 | 35.6333 | 49.9967 | 39.3333 | 35.7287 | 36.1667 | 37.3333 | 41.1167 | 42.45 | 39.3993 | 33.9167 | 40 | 36.5
| Median |  | 36.75 | 38 | 38.8 | 38.05 | 39.15 | 37.05 | 34.98 | 36.0205 | 35.2827 | 35.75 | 50.09 | 39.5 | 35.5183 | 36 | 37.2 | 41.3 | 42.45 | 39.294 | 33.7 | 40 | 36.5
| Std Dev. |  | 0.54191 | 0.5164 | 0.35449 | 0.32042 | 0.54283 | 0.37639 | 0.24967 | 0.49772 | 0.63106 | 0.77632 | 0.41015 | 1.2111 | 1.0069 | 0.40825 | 0.43665 | 1.2432 | 0.18708 | 0.40667 | 0.73869 | 0.63246 | 0.54772
| Rel.Std.Dev. |  | 1.468% | 1.371% | 0.9133% | 0.8425% | 1.393% | 1.018% | 0.7124% | 1.386% | 1.787% | 2.179% | 0.8204% | 3.079% | 2.818% | 1.129% | 1.17% | 3.024% | 0.4407% | 1.032% | 2.178% | 1.581% | 1.501%
| PDM3 |  | -1.72% | 0.2768% | 3.338% | 1.253% | 3.738% | -1.542% | -6.694% | -4.365% | -5.967% | -5.136% | 33.1% | 4.714% | -4.883% | -3.717% | -0.6106% | 9.461% | 13.01% | 4.889% | -9.707% | 6.489% | -2.829%
| Z-Score (Absolute) |  | 0.17 | 0.14 | 0.61 | 0.29 | 0.67 | 0.14 | 0.94 | 0.58 | 0.83 | 0.7 | 5.23 | 0.83 | 0.66 | 0.48 | 0 | 1.56 | 2.11 | 0.85 | 1.41 | 1.1 | 0.34
| Analytical results for Lu in OREAS 505b (Certified Value 0.098 ppm) | [28 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.01: 1.09; 1.04: 1.10; 1.05: 1.11; 1.06: 1.17; 1.07: 1.29 | [6 empty] | ppm
|  |  | 1.01: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES/MS | [6 empty] | 3
|  |  | 1.01: 0.5g; 1.04: 0.5g; 1.05: 0.2g; 1.06: 0.25g; 1.07: 0.25g | [6 empty] | 3
| 1 | 1 | 1.01: 0.095; 1.04: 0.0739; 1.05: 0.10426; 1.06: 0.12; 1.07: 0.1 | [6 empty] | 1
| 1 | 2 | 1.01: 0.097; 1.04: 0.0775; 1.05: 0.10014; 1.06: 0.11; 1.07: 0.1 | [6 empty] | 29
| 1 | 3 | 1.01: 0.095; 1.04: 0.0766; 1.05: 0.10211; 1.06: 0.12; 1.07: 0.09 | [6 empty] | 16
| 1 | 4 | 1.01: 0.094; 1.04: 0.0792; 1.05: 0.097326; 1.06: 0.11; 1.07: 0.1 | [6 empty] | 0.097603
| 1 | 5 | 1.01: 0.098; 1.04: 0.0864; 1.05: 0.098422; 1.06: 0.11; 1.07: 0.1 | [6 empty] | 94
| 1 | 6 | 1.01: 0.098; 1.04: 0.0789; 1.05: 0.096328; 1.06: 0.12; 1.07: 0.1
| Mean |  | 1.01: 0.096167; 1.04: 0.07875; 1.05: 0.099765; 1.06: 0.115; 1.07: 0.098333
| Median |  | 1.01: 0.096; 1.04: 0.0782; 1.05: 0.099279; 1.06: 0.115; 1.07: 0.1
| Std Dev. |  | 1.01: 0.0017224; 1.04: 0.0042061; 1.05: 0.0030144; 1.06: 0.0054772; 1.07: 0.0040825
| Rel.Std.Dev. |  | 1.01: 1.791%; 1.04: 5.341%; 1.05: 3.022%; 1.06: 4.763%; 1.07: 4.152%
| PDM3 |  | 1.01: -1.472%; 1.04: -19.32%; 1.05: 2.215%; 1.06: 17.82%; 1.07: 0.7484%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.04: 6.09; 1.05: 0.45; 1.06: 5.19; 1.07: 0
| Analytical results for Mg in OREAS 505b (Certified Value 0.736 wt.%) | [28 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 3
| 1 | 1 | 0.74 | 0.69 | 0.73 | 0.78 | 0.76 | 0.74 | 0.75 | 0.71 | 0.734 | 0.71432 | 0.73944 | 0.71 | 1.0187 | 0.81149 | 0.74 | 0.71 | 0.71692 | 0.76 | 0.69 | 0.76 | 0.77 | 0.71 | 0.76691 | 0.72 | 0.77 | 0.713 |  | 1
| 1 | 2 | 0.74 | 0.7 | 0.73 | 0.78 | 0.76 | 0.74 | 0.75 | 0.72 | 0.728 | 0.71449 | 0.72177 | 0.69 | 0.8903 | 0.80631 | 0.74 | 0.72 | 0.75143 | 0.75 | 0.7 | 0.78 | 0.78 | 0.71 | 0.77895 | 0.72 | 0.78 | 0.714 |  | #N/A
| 1 | 3 | 0.76 | 0.72 | 0.73 | 0.79 | 0.76 | 0.74 | 0.75 | 0.72 | 0.716 | 0.7239 | 0.71921 | 0.73 | 1.0037 | 0.80845 | 0.75 | 0.73 | 0.73816 | 0.75 | 0.7 | 0.76 | 0.78 | 0.75 | 0.75314 | 0.73 | 0.78 | 0.77 |  | 16
| 1 | 4 | 0.73 | 0.72 | 0.72 | 0.79 | 0.76 | 0.73 | 0.73 | 0.72 | 0.7 | 0.71338 | 0.72666 | 0.7 | 1.0648 | 0.80382 | 0.73 | 0.7 | 0.72804 | 0.75 | 0.69 | 0.73 | 0.77 | 0.74 | 0.75626 | 0.72 | 0.78 | 0.778 |  | 0.73647
| 1 | 5 | 0.74 | 0.73 | 0.73 | 0.79 | 0.75 | 0.73 | 0.72 | 0.72 | 0.72 | 0.71903 | 0.72185 | 0.69 | 1.0649 | 0.80582 | 0.73 | 0.73 | 0.76078 | 0.74 | 0.71 | 0.73 | 0.77 | 0.76 | 0.76728 | 0.71 | 0.78 | 0.713 |  | 95
| 1 | 6 | 0.75 | 0.71 | 0.72 | 0.79 | 0.74 | 0.73 | 0.72 | 0.72 | 0.732 | 0.71733 | 0.71976 | 0.72 | 0.9674 | 0.80477 | 0.74 | 0.71 | 0.7275 | 0.74 | 0.68 | 0.75 | 0.79 | 0.76 | 0.77368 | 0.71 | 0.79 | 0.716
| Mean |  | 0.74333 | 0.71167 | 0.72667 | 0.78667 | 0.755 | 0.735 | 0.73667 | 0.71833 | 0.72167 | 0.71707 | 0.72478 | 0.70667 | 1.0016 | 0.80678 | 0.73833 | 0.71667 | 0.73714 | 0.74833 | 0.695 | 0.75167 | 0.77667 | 0.73833 | 0.76604 | 0.71833 | 0.78 | 0.734
| Median |  | 0.74 | 0.715 | 0.73 | 0.79 | 0.76 | 0.735 | 0.74 | 0.72 | 0.724 | 0.71591 | 0.72181 | 0.705 | 1.0112 | 0.80607 | 0.74 | 0.715 | 0.7331 | 0.75 | 0.695 | 0.755 | 0.775 | 0.745 | 0.7671 | 0.72 | 0.78 | 0.715
| Std Dev. |  | 0.010328 | 0.01472 | 0.005164 | 0.005164 | 0.0083666 | 0.0054772 | 0.015055 | 0.0040825 | 0.012675 | 0.0039591 | 0.0076467 | 0.01633 | 0.066109 | 0.0027899 | 0.0075277 | 0.012111 | 0.016425 | 0.0075277 | 0.010488 | 0.019408 | 0.008165 | 0.023166 | 0.0098938 | 0.0075277 | 0.0063246 | 0.031106
| Rel.Std.Dev. |  | 1.389% | 2.068% | 0.7106% | 0.6564% | 1.108% | 0.7452% | 2.044% | 0.5683% | 1.756% | 0.5521% | 1.055% | 2.311% | 6.6% | 0.3458% | 1.02% | 1.69% | 2.228% | 1.006% | 1.509% | 2.582% | 1.051% | 3.138% | 1.292% | 1.048% | 0.8108% | 4.238%
| PDM3 |  | 0.9322% | -3.368% | -1.331% | 6.816% | 2.516% | -0.1993% | 0.02698% | -2.462% | -2.01% | -2.633% | -1.587% | -4.047% | 36% | 9.547% | 0.2533% | -2.689% | 0.09086% | 1.611% | -5.631% | 2.064% | 5.458% | 0.2533% | 4.015% | -2.462% | 5.911% | -0.3351%
| Z-Score (Absolute) |  | 0.24 | 0.93 | 0.38 | 1.83 | 0.67 | 0.07 | 0.01 | 0.68 | 0.56 | 0.73 | 0.45 | 1.11 | 9.74 | 2.57 | 0.05 | 0.74 | 0.01 | 0.42 | 1.54 | 0.54 | 1.46 | 0.05 | 1.07 | 0.68 | 1.59 | 0.11
| Analytical results for Mn in OREAS 505b (Certified Value 0.038 wt.%) | [28 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  | 3
| 1 | 1 | 0.0386 | 0.0354 | 0.0373 | 0.039 | 0.0392 | 0.0398 | 0.0393 | 0.0366 | 0.0381 | 0.036655 | 0.038045 | 0.0377 | 0.038033 | 0.03728 | 0.0379 | 0.036875 | 0.038 | 0.0368 | 0.0395 | 0.0374 | 0.038 | 0.037925 | 0.0406 | 0.0379 | 0.0366 |  |  | 1
| 1 | 2 | 0.0381 | 0.036 | 0.0369 | 0.039 | 0.0393 | 0.0393 | 0.0393 | 0.0374 | 0.03786 | 0.036471 | 0.036947 | 0.0367 | 0.037085 | 0.03805 | 0.0384 | 0.038599 | 0.0379 | 0.0369 | 0.0406 | 0.0372 | 0.0381 | 0.037936 | 0.0394 | 0.0386 | 0.0364 |  |  | 30
| 1 | 3 | 0.0384 | 0.0365 | 0.0371 | 0.0392 | 0.0393 | 0.0397 | 0.0393 | 0.037 | 0.0373 | 0.036849 | 0.0368 | 0.0389 | 0.03752 | 0.03829 | 0.0381 | 0.036986 | 0.0376 | 0.0365 | 0.0396 | 0.0373 | 0.0396 | 0.038184 | 0.0393 | 0.0377 | 0.0364 |  |  | 16
| 1 | 4 | 0.0386 | 0.0367 | 0.0369 | 0.0394 | 0.0393 | 0.0391 | 0.0382 | 0.0371 | 0.03769 | 0.036446 | 0.037436 | 0.0374 | 0.037708 | 0.03713 | 0.0374 | 0.037618 | 0.0372 | 0.0365 | 0.0376 | 0.0373 | 0.0392 | 0.037823 | 0.0396 | 0.0382 | 0.036 |  |  | 0.037909
| 1 | 5 | 0.0386 | 0.037 | 0.0369 | 0.0397 | 0.0392 | 0.0392 | 0.0375 | 0.0371 | 0.03801 | 0.036771 | 0.036796 | 0.0372 | 0.037054 | 0.03772 | 0.0371 | 0.038952 | 0.0377 | 0.0371 | 0.0374 | 0.0374 | 0.0402 | 0.038479 | 0.0388 | 0.0389 | 0.0366 |  |  | 96
| 1 | 6 | 0.0386 | 0.0365 | 0.0368 | 0.0392 | 0.0387 | 0.0394 | 0.0372 | 0.0376 | 0.03808 | 0.037388 | 0.037056 | 0.0383 | 0.036785 | 0.03767 | 0.0377 | 0.037527 | 0.0385 | 0.0356 | 0.0391 | 0.0377 | 0.0407 | 0.038419 | 0.039 | 0.0387 | 0.0368
| Mean |  | 0.038483 | 0.03635 | 0.036983 | 0.03925 | 0.039167 | 0.039417 | 0.038467 | 0.037133 | 0.03784 | 0.036763 | 0.03718 | 0.0377 | 0.037364 | 0.03769 | 0.037767 | 0.037759 | 0.037817 | 0.036567 | 0.038967 | 0.037383 | 0.0393 | 0.038128 | 0.03945 | 0.038333 | 0.036467
| Median |  | 0.0386 | 0.0365 | 0.0369 | 0.0392 | 0.03925 | 0.03935 | 0.03875 | 0.0371 | 0.037935 | 0.036713 | 0.037001 | 0.03755 | 0.037303 | 0.037695 | 0.0378 | 0.037573 | 0.0378 | 0.03665 | 0.0393 | 0.03735 | 0.0394 | 0.03806 | 0.03935 | 0.0384 | 0.0365
| Std Dev. |  | 0.00020412 | 0.00056833 | 0.00018348 | 0.00026646 | 0.00023381 | 0.00027869 | 0.00096885 | 0.00034448 | 0.00030614 | 0.00034521 | 0.00048474 | 0.00079246 | 0.00046892 | 0.00044104 | 0.00047188 | 0.00084635 | 0.00043551 | 0.00052789 | 0.0012404 | 0.00017224 | 0.0010954 | 0.00027657 | 0.00063166 | 0.0004761 | 0.00027325
| Rel.Std.Dev. |  | 0.5304% | 1.563% | 0.4961% | 0.6789% | 0.597% | 0.707% | 2.519% | 0.9277% | 0.809% | 0.939% | 1.304% | 2.102% | 1.255% | 1.17% | 1.249% | 2.241% | 1.152% | 1.444% | 3.183% | 0.4607% | 2.787% | 0.7254% | 1.601% | 1.242% | 0.7493%
| PDM3 |  | 1.515% | -4.112% | -2.442% | 3.537% | 3.318% | 3.977% | 1.471% | -2.046% | -0.182% | -3.022% | -1.923% | -0.5513% | -1.437% | -0.5776% | -0.3754% | -0.3947% | -0.2435% | -3.541% | 2.79% | -1.387% | 3.669% | 0.5767% | 4.065% | 1.119% | -3.805%
| Z-Score (Absolute) |  | 0.69 | 1.36 | 0.75 | 1.43 | 1.35 | 1.59 | 0.67 | 0.61 | 0.07 | 0.97 | 0.57 | 0.06 | 0.39 | 0.07 | 0 | 0.01 | 0.05 | 1.16 | 1.16 | 0.37 | 1.48 | 0.35 | 1.62 | 0.55 | 1.25
| Analytical results for Mo in OREAS 505b (Certified Value 67 ppm) | [28 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  | 0
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g |  | 1
| 1 | 1 | 68.4 | 55.3 | 67.84 | 72.57 | 70.9 | 64.2 | 61.9 | 70 | 67.59 | 65.3925 | 71.8 | 68.23 | 69.35 | 65.4 | 71.4 | 63.3116 | 69 | 64 | 61.8 | 66.1 | 67.02 | 64.98 | 63.3436 | 66.9 | 70.36 | 72 |  | 1
| 1 | 2 | 68.5 | 55.6 | 68.26 | 70.13 | 71.8 | 63.2 | 60.3 | 68.8 | 67.24 | 64.0347 | 67.3 | 67.38 | 63.92 | 64.9 | 69.1 | 64.3108 | 68 | 67.7 | 64 | 66.9 | 67.82 | 65.07 | 63.3473 | 66 | 72.34 | 70 |  | 31
| 1 | 3 | 69.1 | 59.3 | 67.52 | 70.69 | 71.2 | 65.3 | 62.2 | 69.3 | 66.91 | 63.0495 | 71 | 66.78 | 69.33 | 67.1 | 73 | 62.7474 | 67 | 63.9 | 61.5 | 67.4 | 68.9 | 66.25 | 62.9981 | 66.3 | 70.88 | 72 |  | 16
| 1 | 4 | 69.1 | 57.5 | 68 | 70.18 | 69.7 | 63.4 | 61 | 70.1 | 67.17 | 64.5634 | 71.7 | 67.55 | 68.54 | 66 | 70 | 63.0945 | 67 | 65.7 | 60.3 | 67.1 | 65.74 | 65.25 | 63.3684 | 66.1 | 71.5 | 73 |  | 66.9794
| 1 | 5 | 68.6 | 58.7 | 68.04 | 69.34 | 71.2 | 62.8 | 59.2 | 70.8 | 67.29 | 64.0828 | 69.9 | 66.55 | 72.77 | 65.3 | 71.8 | 64.2484 | 68 | 66.1 | 59.1 | 67.3 | 69.1 | 65.4 | 63.3773 | 65.6 | 71.66 | 65 |  | 97
| 1 | 6 | 69.4 | 58.8 | 67.35 | 69 | 70.2 | 62 | 58.1 | 68.5 | 67.4 | 63.975 | 69.8 | 65.59 | 69.75 | 65.6 | 69.2 | 64.3291 | 67 | 64.3 | 60.8 | 67.7 | 66.23 | 65.12 | 63.3412 | 65 | 71.98 | 65
| Mean |  | 68.85 | 57.5333 | 67.835 | 70.3183 | 70.8333 | 63.4833 | 60.45 | 69.5833 | 67.2667 | 64.183 | 70.25 | 67.0133 | 68.9433 | 65.7167 | 70.75 | 63.6736 | 67.6667 | 65.2833 | 61.25 | 67.0833 | 67.4683 | 65.345 | 63.296 | 65.9833 | 71.4533 | 69.5
| Median |  | 68.85 | 58.1 | 67.92 | 70.155 | 71.05 | 63.3 | 60.65 | 69.65 | 67.265 | 64.0588 | 70.45 | 67.08 | 69.34 | 65.5 | 70.7 | 63.78 | 67.5 | 65 | 61.15 | 67.2 | 67.42 | 65.185 | 63.3454 | 66.05 | 71.58 | 71
| Std Dev. |  | 0.40373 | 1.7212 | 0.34186 | 1.2616 | 0.7607 | 1.1462 | 1.5859 | 0.87044 | 0.22809 | 0.77043 | 1.679 | 0.91627 | 2.8638 | 0.7679 | 1.5668 | 0.70575 | 0.8165 | 1.4972 | 1.6526 | 0.55287 | 1.3826 | 0.46702 | 0.14667 | 0.64317 | 0.72561 | 3.6194
| Rel.Std.Dev. |  | 0.5864% | 2.992% | 0.504% | 1.794% | 1.074% | 1.805% | 2.623% | 1.251% | 0.3391% | 1.2% | 2.39% | 1.367% | 4.154% | 1.168% | 2.215% | 1.108% | 1.207% | 2.293% | 2.698% | 0.8242% | 2.049% | 0.7147% | 0.2317% | 0.9747% | 1.015% | 5.208%
| PDM3 |  | 2.793% | -14.1% | 1.277% | 4.985% | 5.754% | -5.22% | -9.748% | 3.888% | 0.429% | -4.175% | 4.883% | 0.05073% | 2.932% | -1.885% | 5.63% | -4.935% | 1.026% | -2.532% | -8.554% | 0.1552% | 0.73% | -2.44% | -5.499% | -1.487% | 6.68% | 3.763%
| Z-Score (Absolute) |  | 0.48 | 2.75 | 0.19 | 0.9 | 1.04 | 1.05 | 1.92 | 0.69 | 0.03 | 0.85 | 0.88 | 0.05 | 0.5 | 0.42 | 1.02 | 1 | 0.14 | 0.54 | 1.69 | 0.03 | 0.08 | 0.52 | 1.11 | 0.34 | 1.22 | 0.66
| Analytical results for Na in OREAS 505b (Certified Value 0.153 wt.%) | [28 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  | 3
| 1 | 1 | 0.14 | 0.119 | 0.161 | 0.15 | 0.16 | 0.18 | 0.18 | 0.15 | 0.164 | 0.16493 | 0.16137 | 0.12 | 0.13261 | 0.13 | 0.12 | 0.1603 | 0.16 | 0.16 | 0.2 | 0.18 | 0.15 | 0.14926 | 0.15 | 0.14 | 0.147 |  |  | 1
| 1 | 2 | 0.14 | 0.123 | 0.158 | 0.15 | 0.16 | 0.18 | 0.18 | 0.16 | 0.159 | 0.14604 | 0.15943 | 0.12 | 0.11848 | 0.14 | 0.13 | 0.15466 | 0.16 | 0.16 | 0.2 | 0.18 | 0.158 | 0.15306 | 0.15 | 0.13 | 0.15 |  |  | #N/A
| 1 | 3 | 0.14 | 0.13 | 0.159 | 0.15 | 0.16 | 0.18 | 0.18 | 0.15 | 0.156 | 0.15466 | 0.16085 | 0.12 | 0.12222 | 0.14 | 0.13 | 0.15624 | 0.16 | 0.16 | 0.2 | 0.18 | 0.167 | 0.15378 | 0.14 | 0.14 | 0.149 |  |  | 16
| 1 | 4 | 0.15 | 0.131 | 0.16 | 0.15 | 0.16 | 0.18 | 0.2 | 0.16 | 0.153 | 0.1479 | 0.16191 | 0.12 | 0.14287 | 0.13 | 0.13 | 0.15411 | 0.16 | 0.16 | 0.19 | 0.18 | 0.158 | 0.15283 | 0.14 | 0.13 | 0.146 |  |  | 0.15322
| 1 | 5 | 0.15 | 0.129 | 0.159 | 0.16 | 0.16 | 0.17 | 0.2 | 0.16 | 0.158 | 0.14842 | 0.15891 | 0.12 | 0.12797 | 0.13 | 0.13 | 0.1595 | 0.16 | 0.16 | 0.18 | 0.18 | 0.164 | 0.15647 | 0.14 | 0.14 | 0.146 |  |  | 98
| 1 | 6 | 0.15 | 0.126 | 0.16 | 0.16 | 0.15 | 0.17 | 0.2 | 0.16 | 0.16 | 0.15553 | 0.15894 | 0.12 | 0.11981 | 0.14 | 0.13 | 0.15516 | 0.16 | 0.15 | 0.2 | 0.18 | 0.162 | 0.15877 | 0.14 | 0.14 | 0.146
| Mean |  | 0.145 | 0.12633 | 0.1595 | 0.15333 | 0.15833 | 0.17667 | 0.19 | 0.15667 | 0.15833 | 0.15291 | 0.16024 | 0.12 | 0.12733 | 0.135 | 0.12833 | 0.15666 | 0.16 | 0.15833 | 0.195 | 0.18 | 0.15983 | 0.15403 | 0.14333 | 0.13667 | 0.14733
| Median |  | 0.145 | 0.1275 | 0.1595 | 0.15 | 0.16 | 0.18 | 0.19 | 0.16 | 0.1585 | 0.15154 | 0.16014 | 0.12 | 0.1251 | 0.135 | 0.13 | 0.1557 | 0.16 | 0.16 | 0.2 | 0.18 | 0.16 | 0.15342 | 0.14 | 0.14 | 0.1465
| Std Dev. |  | 0.0054772 | 0.0046332 | 0.0010488 | 0.005164 | 0.0040825 | 0.005164 | 0.010954 | 0.005164 | 0.0037238 | 0.0070275 | 0.0013067 | 0 | 0.0092794 | 0.0054772 | 0.0040825 | 0.0026162 | 0 | 0.0040825 | 0.0083666 | 3.0405e-17 | 0.005947 | 0.0032735 | 0.005164 | 0.005164 | 0.0017512
| Rel.Std.Dev. |  | 3.777% | 3.667% | 0.6576% | 3.368% | 2.578% | 2.923% | 5.766% | 3.296% | 2.352% | 4.596% | 0.8155% | 0% | 7.288% | 4.057% | 3.181% | 1.67% | 0% | 2.578% | 4.291% | 1.689e-14% | 3.721% | 2.125% | 3.603% | 3.779% | 1.189%
| PDM3 |  | -5.368% | -17.55% | 4.096% | 0.07096% | 3.334% | 15.3% | 24% | 2.246% | 3.334% | -0.2041% | 4.577% | -21.68% | -16.9% | -11.89% | -16.24% | 2.243% | 4.422% | 3.334% | 27.26% | 17.47% | 4.313% | 0.5239% | -6.455% | -10.81% | -3.845%
| Z-Score (Absolute) |  | 0.84 | 2.19 | 0.21 | 0.24 | 0.12 | 1.45 | 2.41 | 0 | 0.12 | 0.27 | 0.26 | 2.65 | 2.12 | 1.57 | 2.05 | 0 | 0.24 | 0.12 | 2.77 | 1.69 | 0.23 | 0.19 | 0.96 | 1.45 | 0.67
| Analytical results for Nb in OREAS 505b (Indicative Value 0.89 ppm) | [28 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.29 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 01g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | [10 empty] | 2
| 1 | 1 | 1.1 | 0.63 | 0.88 | 0.65 | 0.66 | 0.61 | < 0.1 | 18.874 | 2.06 | 0.92 | 1.8937 | 1.04 | 0.78 | 1.26 | < 5 | 0.6 | 1.02 | [10 empty] | 1
| 1 | 2 | 1.1 | 0.66 | 0.9 | 0.63 | 0.59 | 0.57 | < 0.1 | 18.4926 | 1.33 | 0.91 | 1.804 | 1.02 | 0.78 | 1.24 | < 5 | 0.5 | 1.1 | [10 empty] | 6
| 1 | 3 | 1.2 | 0.66 | 0.93 | 0.72 | 0.61 | 0.61 | < 0.1 | 18.7781 | 1.44 | 0.93 | 1.7888 | 1.04 | 0.66 | 1.18 | < 5 | 0.6 | 1.09 | [10 empty] | 16
| 1 | 4 | 1.3 | 0.63 | 0.88 | 0.86 | 0.64 | 0.57 | < 0.1 | 18.6791 | 1.34 | 0.79 | 2.2054 | 1.06 | 0.68 | 1.24 | < 5 | 0.6 | 1.05 | [10 empty] | 0.88997
| 1 | 5 | 1.2 | 0.66 | 0.89 | 0.85 | 0.55 | 0.6 | < 0.1 | 18.2981 | 1.52 | 0.91 | 2.5373 | 0.97 | 0.67 | 1.02 | < 5 | 0.6 | 1.13 | [10 empty] | 12
| 1 | 6 | 1.1 | 0.74 | 0.78 | 0.72 | 0.56 | 0.68 | < 0.1 | 18.2933 | 1.46 | 0.87 | 2.4989 | 1.01 | 0.65 | 1.18 | < 5 | 0.6 | 1.11
| Mean |  | 1.1667 | 0.66333 | 0.87667 | 0.73833 | 0.60167 | 0.60667 |  | 18.5692 | 1.525 | 0.88833 | 2.1214 | 1.0233 | 0.70333 | 1.1867 |  | 0.58333 | 1.0833
| Median |  | 1.15 | 0.66 | 0.885 | 0.72 | 0.6 | 0.605 |  | 18.5858 | 1.45 | 0.91 | 2.0496 | 1.03 | 0.675 | 1.21 |  | 0.6 | 1.095
| Std Dev. |  | 0.08165 | 0.040332 | 0.050859 | 0.097451 | 0.043551 | 0.040332 |  | 0.24661 | 0.27201 | 0.052313 | 0.34225 | 0.031411 | 0.060222 | 0.088242 |  | 0.040825 | 0.040825
| Rel.Std.Dev. |  | 6.999% | 6.08% | 5.801% | 13.2% | 7.238% | 6.648% |  | 1.328% | 17.84% | 5.889% | 16.13% | 3.07% | 8.562% | 7.436% |  | 6.999% | 3.768%
| PDM3 |  | 31.09% | -25.47% | -1.495% | -17.04% | -32.39% | -31.83% |  | 1986% | 71.35% | -0.1839% | 138.4% | 14.99% | -20.97% | 33.34% |  | -34.45% | 21.73%
| Z-Score (Absolute) |  | NA | 0.54 | 0.03 | 0.36 | 0.69 | 0.67 | 2.01 | 42.33 | 1.52 | 0 | 2.95 | 0.32 | 0.44 | 0.71 | 3.86 | NA | 0.47
Results from laboratories 1.01 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 505b (Indicative Value 15 ppm) | [28 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.05: Lab | [10 empty] | 1
| No | No | 1.01: 1.09; 1.05: 1.10 | [10 empty] | ppm
|  |  | 1.01: AR*MS; 1.05: AR*MS | [10 empty] | 1
|  |  | 1.01: 0.5g; 1.05: 0.5g | [10 empty] | 1
| 1 | 1 | 1.01: 14.955; 1.05: 14.0518 | [10 empty] | 1
| 1 | 2 | 1.01: 14.627; 1.05: 14.568 | [10 empty] | 7
| 1 | 3 | 1.01: 14.954; 1.05: 14.5898 | [10 empty] | 16
| 1 | 4 | 1.01: 14.82; 1.05: 14.9074 | [10 empty] | 15.0472
| 1 | 5 | 1.01: 14.86; 1.05: 16.4086 | [10 empty] | 13
| 1 | 6 | 1.01: 15.169; 1.05: 16.6556
| Mean |  | 1.01: 14.8975; 1.05: 15.1969
| Median |  | 1.01: 14.907; 1.05: 14.7486
| Std Dev. |  | 1.01: 0.17933; 1.05: 1.0729
| Rel.Std.Dev. |  | 1.01: 1.204%; 1.05: 7.06%
| PDM3 |  | 1.01: -0.9948%; 1.05: 0.9948%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.05: 0.67
| Analytical results for Ni in OREAS 505b (Certified Value 38.6 ppm) | [28 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 1
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 2
| 1 | 1 | 36 | 33.9 | 41.5 | 39.8 | 42.1 | 39.7 | 40.5 | 38.5 | 39.4 | 40.0111 | 38.539 | 38.1 | 39.34 | 37.32 | 39.23 | 40.1 | 36.8646 | 38.8 | 37.8 | 38.4 | 38 | 34.2 | 38.159 | 36.6 | 39.1 | 38 |  | 1
| 1 | 2 | 36 | 34.4 | 39.1 | 38.9 | 42.2 | 40.7 | 38.3 | 38 | 39.1 | 39.8819 | 37.8352 | 39.6 | 39.77 | 37.76 | 39.5 | 38.6 | 37.8345 | 37.6 | 38.3 | 39.5 | 38.5 | 34.2 | 37.119 | 35.8 | 39.7 | 38 |  | 17
| 1 | 3 | 36 | 34.7 | 40.1 | 40.1 | 43.6 | 41.3 | 39.7 | 38.2 | 39 | 39.9478 | 37.3425 | 40.8 | 39.01 | 37.52 | 38.84 | 39.7 | 37.1881 | 39.3 | 38.2 | 38.5 | 38.7 | 35.6 | 36.623 | 35.8 | 38.9 | 37 |  | 16
| 1 | 4 | 37 | 35.1 | 40.7 | 39.4 | 42.2 | 40.6 | 42.7 | 38.2 | 39.2 | 40.0309 | 37.9671 | 40.8 | 39.18 | 37.57 | 39.26 | 39.4 | 39.2296 | 38.6 | 36.6 | 39.1 | 38.7 | 36 | 37.122 | 35.7 | 39.2 | 37 |  | 38.5838
| 1 | 5 | 37 | 35.4 | 39.8 | 38.8 | 44.2 | 40.3 | 40.7 | 38.8 | 39.8 | 40.5055 | 37.7899 | 39.2 | 38.989 | 38.54 | 38.32 | 38.7 | 38.3311 | 39.2 | 40.3 | 37.2 | 38.7 | 36.6 | 38.051 | 35.5 | 40.3 | 37 |  | 99
| 1 | 6 | 38 | 34.8 | 41.1 | 37.3 | 44 | 39.3 | 40.6 | 36.1 | 39.8 | 39.8621 | 37.4571 | 38.9 | 39.44 | 37.99 | 39.78 | 41.4 | 38.807 | 40 | 38.2 | 38.1 | 39 | 36.6 | 38.167 | 35.7 | 40.6 | 37
| Mean |  | 36.6667 | 34.7167 | 40.3833 | 39.05 | 43.05 | 40.3167 | 40.4167 | 37.9667 | 39.3833 | 40.0399 | 37.8218 | 39.5667 | 39.2882 | 37.7833 | 39.155 | 39.65 | 38.0425 | 38.9167 | 38.2333 | 38.4667 | 38.6 | 35.5333 | 37.5402 | 35.85 | 39.6333 | 37.3333
| Median |  | 36.5 | 34.75 | 40.4 | 39.15 | 42.9 | 40.45 | 40.55 | 38.2 | 39.3 | 39.9794 | 37.8125 | 39.4 | 39.26 | 37.665 | 39.245 | 39.55 | 38.0828 | 39 | 38.2 | 38.45 | 38.7 | 35.8 | 37.5865 | 35.75 | 39.45 | 37
| Std Dev. |  | 0.8165 | 0.52694 | 0.88638 | 0.99348 | 0.98742 | 0.72226 | 1.4372 | 0.95638 | 0.34881 | 0.23782 | 0.42411 | 1.0746 | 0.29546 | 0.43482 | 0.51434 | 1.033 | 0.92062 | 0.80602 | 1.1944 | 0.80166 | 0.33466 | 1.1003 | 0.66787 | 0.38341 | 0.69186 | 0.5164
| Rel.Std.Dev. |  | 2.227% | 1.518% | 2.195% | 2.544% | 2.294% | 1.791% | 3.556% | 2.519% | 0.8857% | 0.594% | 1.121% | 2.716% | 0.752% | 1.151% | 1.314% | 2.605% | 2.42% | 2.071% | 3.124% | 2.084% | 0.867% | 3.097% | 1.779% | 1.069% | 1.746% | 1.383%
| PDM3 |  | -4.969% | -10.02% | 4.664% | 1.208% | 11.58% | 4.491% | 4.75% | -1.599% | 2.072% | 3.774% | -1.975% | 2.547% | 1.826% | -2.075% | 1.48% | 2.763% | -1.403% | 0.8628% | -0.9083% | -0.3035% | 0.04203% | -7.906% | -2.705% | -7.085% | 2.72% | -3.241%
| Z-Score (Absolute) |  | 1.54 | 2.98 | 1.2 | 0.22 | 3.17 | 1.15 | 1.22 | 0.58 | 0.46 | 0.95 | 0.69 | 0.6 | 0.39 | 0.72 | 0.29 | 0.66 | 0.53 | 0.12 | 0.39 | 0.22 | 0.12 | 2.38 | 0.9 | 2.15 | 0.65 | 1.05
| Analytical results for P in OREAS 505b (Certified Value 0.068 wt.%) | [28 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  |  | 3
| 1 | 1 | 0.0682 | 0.06 | 0.064 | 0.07 | 0.071 | 0.07 | 0.07 | 0.068 | 0.0638 | 0.070226 | 0.068206 | 0.063 | 0.07 | 0.07 | 0.068445 | 0.0692 | 0.067 | 0.075 | 0.0676 | 0.063 | 0.067828 | 0.066 | 0.0712 | 0.065 |  |  |  | 1
| 1 | 2 | 0.0671 | 0.062 | 0.065 | 0.07 | 0.07 | 0.069 | 0.069 | 0.069 | 0.0638 | 0.073709 | 0.067178 | 0.067 | 0.07 | 0.07 | 0.067 | 0.0672 | 0.068 | 0.076 | 0.068 | 0.065 | 0.069224 | 0.064 | 0.0716 | 0.066 |  |  |  | 18
| 1 | 3 | 0.0674 | 0.062 | 0.064 | 0.07 | 0.07 | 0.07 | 0.07 | 0.068 | 0.0651 | 0.073068 | 0.066003 | 0.069 | 0.07 | 0.07 | 0.067734 | 0.0678 | 0.068 | 0.075 | 0.0684 | 0.063 | 0.065691 | 0.064 | 0.0733 | 0.063 |  |  |  | 16
| 1 | 4 | 0.0684 | 0.063 | 0.065 | 0.07 | 0.07 | 0.07 | 0.069 | 0.069 | 0.0645 | 0.071562 | 0.067369 | 0.07 | 0.07 | 0.07 | 0.066939 | 0.0672 | 0.067 | 0.072 | 0.0675 | 0.065 | 0.066888 | 0.064 | 0.0716 | 0.066 |  |  |  | 0.067924
| 1 | 5 | 0.0691 | 0.063 | 0.063 | 0.07 | 0.07 | 0.069 | 0.069 | 0.068 | 0.0656 | 0.069901 | 0.066963 | 0.065 | 0.07 | 0.07 | 0.067997 | 0.0679 | 0.068 | 0.071 | 0.0676 | 0.068 | 0.067109 | 0.063 | 0.0712 | 0.066 |  |  |  | 100
| 1 | 6 | 0.0687 | 0.062 | 0.065 | 0.07 | 0.07 | 0.069 | 0.068 | 0.069 | 0.065 | 0.069425 | 0.066862 | 0.066 | 0.07 | 0.07 | 0.068349 | 0.0667 | 0.065 | 0.073 | 0.0686 | 0.067 | 0.068025 | 0.064 | 0.0723 | 0.066
| Mean |  | 0.06815 | 0.062 | 0.064333 | 0.07 | 0.070167 | 0.0695 | 0.069167 | 0.0685 | 0.064633 | 0.071315 | 0.067097 | 0.066667 | 0.07 | 0.07 | 0.067744 | 0.067667 | 0.067167 | 0.073667 | 0.06795 | 0.065167 | 0.067461 | 0.064167 | 0.071867 | 0.065333
| Median |  | 0.0683 | 0.062 | 0.0645 | 0.07 | 0.07 | 0.0695 | 0.069 | 0.0685 | 0.06475 | 0.070894 | 0.06707 | 0.0665 | 0.07 | 0.07 | 0.067865 | 0.0675 | 0.0675 | 0.074 | 0.0678 | 0.065 | 0.067469 | 0.064 | 0.0716 | 0.066
| Std Dev. |  | 0.00076616 | 0.0010954 | 0.0008165 | 0 | 0.00040825 | 0.00054772 | 0.00075277 | 0.00054772 | 0.00073394 | 0.0017678 | 0.00071846 | 0.002582 | 0 | 0 | 0.00065182 | 0.00087101 | 0.001169 | 0.0019664 | 0.00046368 | 0.0020412 | 0.0011955 | 0.00098319 | 0.00080911 | 0.0012111
| Rel.Std.Dev. |  | 1.124% | 1.767% | 1.269% | 0% | 0.5818% | 0.7881% | 1.088% | 0.7996% | 1.136% | 2.479% | 1.071% | 3.873% | 0% | 0% | 0.9622% | 1.287% | 1.741% | 2.669% | 0.6824% | 3.132% | 1.772% | 1.532% | 1.126% | 1.854%
| PDM3 |  | 0.3323% | -8.722% | -5.287% | 3.056% | 3.301% | 2.32% | 1.829% | 0.8475% | -4.845% | 4.992% | -1.219% | -1.852% | 3.056% | 3.056% | -0.2657% | -0.3793% | -1.115% | 8.454% | 0.03782% | -4.06% | -0.6821% | -5.532% | 5.804% | -3.815%
| Z-Score (Absolute) |  | 0.09 | 1.83 | 1.1 | 0.67 | 0.73 | 0.52 | 0.41 | 0.2 | 1.01 | 1.09 | 0.23 | 0.37 | 0.67 | 0.67 | 0.03 | 0.06 | 0.21 | 1.82 | 0.03 | 0.84 | 0.12 | 1.15 | 1.26 | 0.79
| Analytical results for Pb in OREAS 505b (Certified Value 34.3 ppm) | [28 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | 1
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 2
| 1 | 1 | 38.8 | 30.1 | 32.12 | 37.1 | 35.4 | 35 | 35 | 35.2 | 34.1 | 28.7116 | 33.9646 | 34.1 | 32.33 | 35.5 | 35.9 | 30.4 | 36.0703 | 32.9 | 33.6 | 34.2 | 33.5 | 36.41 | 33.328 | 29.9967 | 36.9 | 34.2 | 33 | 1
| 1 | 2 | 38.2 | 30.7 | 31.73 | 36.2 | 34.6 | 36.4 | 32.7 | 35.3 | 34.1 | 29.1081 | 33.1296 | 33.9 | 28.77 | 34.83 | 35.9 | 31.4 | 36.3013 | 31.9 | 33.2 | 35.6 | 33.9 | 36.46 | 34.715 | 29.2505 | 36 | 35.4 | 34 | 19
| 1 | 3 | 37.7 | 30.7 | 32.06 | 37.1 | 34.8 | 36 | 35.2 | 35.1 | 34 | 48.3095 | 33.1249 | 35 | 28.04 | 34.86 | 35 | 31.8 | 32.9949 | 33.6 | 33 | 35.2 | 34.1 | 36.43 | 34.693 | 30.6317 | 35.2 | 34.4 | 33 | 16
| 1 | 4 | 39 | 31 | 31.7 | 36.3 | 34.1 | 36.1 | 35.3 | 35.1 | 33.7 | 37.438 | 33.5819 | 34.1 | 38.126 | 35.9 | 34.8 | 30.4 | 33.1672 | 33.4 | 33.6 | 34.5 | 34.2 | 35.78 | 33.906 | 30.1018 | 35.6 | 34.9 | 32 | 34.2693
| 1 | 5 | 38.1 | 31.5 | 30.77 | 35.4 | 35.3 | 35.8 | 34.9 | 35.6 | 34.1 | 34.3033 | 33.4583 | 34.4 | 38.809 | 35.54 | 35.5 | 31.4 | 34.2298 | 33 | 34.5 | 33.4 | 34.2 | 36.77 | 34.713 | 29.9248 | 36.8 | 35.8 | 34 | 101
| 1 | 6 | 38.5 | 30.8 | 31.24 | 34.9 | 34.3 | 35.4 | 34.5 | 34 | 34.8 | 30.3124 | 32.9047 | 34.3 | 28.17 | 36.02 | 34.6 | 30.3 | 36.3326 | 33.8 | 34 | 34.1 | 34.2 | 36.28 | 35.659 | 30.1779 | 36.8 | 35.9 | 33
| Mean |  | 38.3833 | 30.8 | 31.6033 | 36.1667 | 34.75 | 35.7833 | 34.6 | 35.05 | 34.1333 | 34.6972 | 33.3607 | 34.3 | 32.3742 | 35.4417 | 35.2833 | 30.95 | 34.8494 | 33.1 | 33.65 | 34.5 | 34.0167 | 36.355 | 34.5023 | 30.0139 | 36.2167 | 35.1 | 33.1667
| Median |  | 38.35 | 30.75 | 31.715 | 36.25 | 34.7 | 35.9 | 34.95 | 35.15 | 34.1 | 32.3079 | 33.2939 | 34.2 | 30.55 | 35.52 | 35.25 | 30.9 | 35.1501 | 33.2 | 33.6 | 34.35 | 34.15 | 36.42 | 34.703 | 30.0493 | 36.4 | 35.15 | 33
| Std Dev. |  | 0.47924 | 0.45607 | 0.51523 | 0.88919 | 0.5244 | 0.50761 | 0.9716 | 0.54681 | 0.36148 | 7.4733 | 0.38474 | 0.38471 | 4.9788 | 0.504 | 0.56362 | 0.65651 | 1.5781 | 0.68118 | 0.54314 | 0.79498 | 0.27869 | 0.32501 | 0.79981 | 0.44885 | 0.72226 | 0.71554 | 0.75277
| Rel.Std.Dev. |  | 1.249% | 1.481% | 1.63% | 2.459% | 1.509% | 1.419% | 2.808% | 1.56% | 1.059% | 21.54% | 1.153% | 1.122% | 15.38% | 1.422% | 1.597% | 2.121% | 4.528% | 2.058% | 1.614% | 2.304% | 0.8193% | 0.894% | 2.318% | 1.495% | 1.994% | 2.039% | 2.27%
| PDM3 |  | 12.01% | -10.12% | -7.779% | 5.537% | 1.403% | 4.418% | 0.9651% | 2.278% | -0.3967% | 1.249% | -2.651% | 0.08965% | -5.53% | 3.421% | 2.959% | -9.686% | 1.693% | -3.412% | -1.807% | 0.6733% | -0.7371% | 6.086% | 0.6801% | -12.42% | 5.683% | 2.424% | -3.217%
| Z-Score (Absolute) |  | 2.79 | 2.66 | 2.08 | 1.19 | 0.18 | 0.92 | 0.07 | 0.39 | 0.26 | 0.14 | 0.82 | 0.15 | 1.53 | 0.67 | 0.56 | 2.55 | 0.25 | 1.01 | 0.61 | 0 | 0.35 | 1.33 | 0 | 3.22 | 1.23 | 0.43 | 0.96
| Analytical results for Pd in OREAS 505b (Indicative Value 22.7 ppb) | [28 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.18 | ppb
|  |  | 1.01: AR*MS; 1.02: AR*OES/MS | 1
|  |  | 1.01: 0.5g; 1.02: 01g | 1
| 1 | 1 | 1.01: 23; 1.02: 21.0486 | 1
| 1 | 2 | 1.01: 26; 1.02: 21.4284 | 1
| 1 | 3 | 1.01: 24; 1.02: 21.755 | 16
| 1 | 4 | 1.01: 24; 1.02: 22.4233 | 22.6527
| 1 | 5 | 1.01: 22; 1.02: 21.1135 | 14
| 1 | 6 | 1.01: 23; 1.02: 22.0634
| Mean |  | 1.01: 23.6667; 1.02: 21.6387
| Median |  | 1.01: 23.5; 1.02: 21.5917
| Std Dev. |  | 1.01: 1.3663; 1.02: 0.54365
| Rel.Std.Dev. |  | 1.01: 5.773%; 1.02: 2.512%
| PDM3 |  | 1.01: 4.476%; 1.02: -4.476%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Pr in OREAS 505b (Indicative Value 3.76 ppm) | [28 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g | 2
| 1 | 1 | 1.01: 3.79; 1.02: 3.52; 1.03: 5.4103 | 1
| 1 | 2 | 1.01: 3.732; 1.02: 3.5803; 1.03: 5.3699 | 9
| 1 | 3 | 1.01: 3.776; 1.02: 3.6913; 1.03: 5.3903 | 16
| 1 | 4 | 1.01: 3.788; 1.02: 3.4847; 1.03: 5.3437 | 3.7619
| 1 | 5 | 1.01: 3.806; 1.02: 3.8879; 1.03: 5.131 | 15
| 1 | 6 | 1.01: 3.867; 1.02: 4.2193; 1.03: 5.2313
| Mean |  | 1.01: 3.7932; 1.02: 3.7306; 1.03: 5.3128
| Median |  | 1.01: 3.789; 1.02: 3.6358; 1.03: 5.3568
| Std Dev. |  | 1.01: 0.04402; 1.02: 0.28019; 1.03: 0.109
| Rel.Std.Dev. |  | 1.01: 1.161%; 1.02: 7.511%; 1.03: 2.052%
| PDM3 |  | 1.01: 0.8318%; 1.02: -0.8318%; 1.03: 41.23%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 16.37
| Analytical results for Pt in OREAS 505b (Indicative Value 1.33 ppb) | [28 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.02: Lab | 1
| No | No | 1.01: 1.09; 1.02: 1.18 | ppb
|  |  | 1.01: AR*MS; 1.02: AR*OES/MS | 2
|  |  | 1.01: 0.5g; 1.02: 01g | 2
| 1 | 1 | 1.01: 2; 1.02: < 5 | 1
| 1 | 2 | 1.01: 1; 1.02: < 5 | 1
| 1 | 3 | 1.01: 1; 1.02: < 5 | 16
| 1 | 4 | 1.01: 1; 1.02: < 5 | 1.3333
| 1 | 5 | 1.01: 1; 1.02: < 5 | 16
| 1 | 6 | 1.01: 2; 1.02: < 5
| Mean |  | 1.01: 1.3333
Median | bar: 1.01
| Std Dev. |  | 1.01: 0.5164
| Rel.Std.Dev. |  | 1.01: 38.73%
| PDM3 |  | 1.01: 2.442e-13%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Rb in OREAS 505b (Certified Value 89 ppm) | [28 empty] | Certified
Rb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.16: Lab; 1.17: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.04; 1.03: 1.05; 1.04: 1.06; 1.05: 1.07; 1.06: 1.08; 1.07: 1.09; 1.08: 1.11; 1.09: 1.12; 1.10: 1.17; 1.11: 1.18; 1.12: 1.19; 1.13: 1.20; 1.14: 1.21; 1.16: 1.27; 1.17: 1.29 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*MS; 1.08: AR*MS; 1.09: AR*OES/MS; 1.10: AR*OES/MS; 1.11: AR*OES/MS; 1.12: AR*MS; 1.13: AR*MS; 1.14: AR*MS; 1.16: AR*OES/MS; 1.17: AR*OES/MS | 0
|  |  | 1.01: 01g; 1.02: 0.25g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.2g; 1.09: 0.5g; 1.10: 0.25g; 1.11: 01g; 1.12: 0.2g; 1.13: 0.5g; 1.14: 0.5g; 1.16: 0.25g; 1.17: 0.25g | 1
| 1 | 1 | 1.01: 89.2; 1.02: 90.9; 1.03: 88.2; 1.04: 90.5; 1.05: 90.2; 1.06: 90.3; 1.07: 87.36; 1.08: 88.5572; 1.09: 91.5; 1.10: 88.6; 1.11: 79.7927; 1.12: 79.9; 1.13: 87.7; 1.14: 90.4; 1.16: 86.2; 1.17: 97.1 | 1
| 1 | 2 | 1.01: 89.6; 1.02: 89.6; 1.03: 87.7; 1.04: 90.6; 1.05: 87.6; 1.06: 88.1; 1.07: 88.47; 1.08: 86.1176; 1.09: 83.7; 1.10: 88.6; 1.11: 83.178; 1.12: 74.8; 1.13: 89.1; 1.14: 92.2; 1.16: 84.8; 1.17: 98.1 | 37
| 1 | 3 | 1.01: 88.9; 1.02: 91.4; 1.03: 92.7; 1.04: 92.5; 1.05: 90; 1.06: 90.2; 1.07: 89.31; 1.08: 86.3591; 1.09: 84.4; 1.10: 87.3; 1.11: 79.6163; 1.12: 79.2; 1.13: 89.3; 1.14: 90.2; 1.16: 83.8; 1.17: 96.9 | 16
| 1 | 4 | 1.01: 88.1; 1.02: 89.9; 1.03: 88.3; 1.04: 89.9; 1.05: 93.7; 1.06: 90.5; 1.07: 87.71; 1.08: 88.906; 1.09: 86.6; 1.10: 86.4; 1.11: 80.5958; 1.12: 78.3; 1.13: 87.6; 1.14: 87.1; 1.16: 84.4; 1.17: 98.2 | 89.3042
| 1 | 5 | 1.01: 88.4; 1.02: 89.6; 1.03: 90.5; 1.04: 89.9; 1.05: 90.9; 1.06: 90; 1.07: 91.32; 1.08: 85.1879; 1.09: 84.2; 1.10: 88.8; 1.11: 83.8575; 1.12: 79.8; 1.13: 91.7; 1.14: 86.6; 1.16: 83.1; 1.17: 96.6 | 102
| 1 | 6 | 1.01: 88.8; 1.02: 88.1; 1.03: 94.2; 1.04: 88.1; 1.05: 89.8; 1.06: 84.7; 1.07: 89.63; 1.08: 85.0325; 1.09: 85.9; 1.10: 89.5; 1.11: 83.0422; 1.12: 81.6; 1.13: 91.5; 1.14: 90.8; 1.16: 83.2; 1.17: 101.8
| Mean |  | 1.01: 88.8333; 1.02: 89.9167; 1.03: 90.2667; 1.04: 90.25; 1.05: 90.3667; 1.06: 88.9667; 1.07: 88.9667; 1.08: 86.6934; 1.09: 86.05; 1.10: 88.2; 1.11: 81.6804; 1.12: 78.9333; 1.13: 89.4833; 1.14: 89.55; 1.16: 84.25; 1.17: 98.1167
| Median |  | 1.01: 88.85; 1.02: 89.75; 1.03: 89.4; 1.04: 90.2; 1.05: 90.1; 1.06: 90.1; 1.07: 88.89; 1.08: 86.2384; 1.09: 85.15; 1.10: 88.6; 1.11: 81.819; 1.12: 79.5; 1.13: 89.2; 1.14: 90.3; 1.16: 84.1; 1.17: 97.6
| Std Dev. |  | 1.01: 0.53914; 1.02: 1.1548; 1.03: 2.6897; 1.04: 1.4223; 1.05: 1.9765; 1.06: 2.266; 1.07: 1.4494; 1.08: 1.6636; 1.09: 2.8877; 1.10: 1.1331; 1.11: 1.8887; 1.12: 2.2958; 1.13: 1.7826; 1.14: 2.2107; 1.16: 1.1623; 1.17: 1.9177
| Rel.Std.Dev. |  | 1.01: 0.6069%; 1.02: 1.284%; 1.03: 2.98%; 1.04: 1.576%; 1.05: 2.187%; 1.06: 2.547%; 1.07: 1.629%; 1.08: 1.919%; 1.09: 3.356%; 1.10: 1.285%; 1.11: 2.312%; 1.12: 2.909%; 1.13: 1.992%; 1.14: 2.469%; 1.16: 1.38%; 1.17: 1.955%
| PDM3 |  | 1.01: -0.5273%; 1.02: 0.6858%; 1.03: 1.078%; 1.04: 1.059%; 1.05: 1.19%; 1.06: -0.378%; 1.07: -0.378%; 1.08: -2.924%; 1.09: -3.644%; 1.10: -1.236%; 1.11: -8.537%; 1.12: -11.61%; 1.13: 0.2005%; 1.14: 0.2752%; 1.16: -5.66%; 1.17: 9.868%
| Z-Score (Absolute) |  | 1.01: 0.07; 1.02: 0.5; 1.03: 0.68; 1.04: 0.67; 1.05: 0.73; 1.06: 0; 1.07: 0; 1.08: 1.19; 1.09: 1.52; 1.10: 0.4; 1.11: 3.8; 1.12: 5.24; 1.13: 0.27; 1.14: 0.3; 1.16: 2.46; 1.17: 4.78
| Analytical results for Re in OREAS 505b (Certified Value 0.022 ppm) | [28 empty] | Certified
Re | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.06; 1.04: 1.07; 1.05: 1.08; 1.06: 1.09; 1.07: 1.12; 1.08: 1.17; 1.09: 1.18; 1.10: 1.19; 1.11: 1.20; 1.12: 1.21; 1.13: 1.27; 1.14: 1.29 | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS | 3
|  |  | 1.01: 01g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 01g; 1.10: 0.2g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.25g; 1.14: 0.25g | 3
| 1 | 1 | 1.01: < 0.05; 1.02: 0.023; 1.03: 0.021; 1.04: 0.022; 1.05: 0.022; 1.06: 0.022; 1.07: 0.129; 1.08: 0.03; 1.09: < 0.05; 1.10: 0.023; 1.11: 0.022; 1.12: 0.021; 1.13: 0.022; 1.14: 0.021 | 1
| 1 | 2 | 1.01: < 0.05; 1.02: 0.024; 1.03: 0.022; 1.04: 0.021; 1.05: 0.021; 1.06: 0.023; 1.07: 0.021; 1.08: 0.03; 1.09: < 0.05; 1.10: 0.022; 1.11: 0.019; 1.12: 0.022; 1.13: 0.022; 1.14: 0.021 | 38
| 1 | 3 | 1.01: < 0.05; 1.02: 0.023; 1.03: 0.021; 1.04: 0.021; 1.05: 0.021; 1.06: 0.023; 1.07: 0.019; 1.08: 0.03; 1.09: < 0.05; 1.10: 0.022; 1.11: 0.021; 1.12: 0.022; 1.13: 0.02; 1.14: 0.022 | 16
| 1 | 4 | 1.01: < 0.05; 1.02: 0.021; 1.03: 0.021; 1.04: 0.023; 1.05: 0.022; 1.06: 0.022; 1.07: 0.022; 1.08: 0.03; 1.09: < 0.05; 1.10: 0.023; 1.11: 0.021; 1.12: 0.022; 1.13: 0.023; 1.14: 0.021 | 0.021811
| 1 | 5 | 1.01: < 0.05; 1.02: 0.024; 1.03: 0.022; 1.04: 0.021; 1.05: 0.026; 1.06: 0.024; 1.07: 0.027; 1.08: 0.03; 1.09: < 0.05; 1.10: 0.022; 1.11: 0.021; 1.12: 0.022; 1.13: 0.022; 1.14: 0.023 | 103
| 1 | 6 | 1.01: < 0.05; 1.02: 0.022; 1.03: 0.021; 1.04: 0.019; 1.05: 0.024; 1.06: 0.023; 1.07: 0.021; 1.08: 0.03; 1.09: < 0.05; 1.10: 0.023; 1.11: 0.022; 1.12: 0.021; 1.13: 0.021; 1.14: 0.021
| Mean |  | 1.02: 0.022833; 1.03: 0.021333; 1.04: 0.021167; 1.05: 0.022667; 1.06: 0.022833; 1.07: 0.039833; 1.08: 0.03; 1.10: 0.0225; 1.11: 0.021; 1.12: 0.021667; 1.13: 0.021667; 1.14: 0.0215
| Median |  | 1.02: 0.023; 1.03: 0.021; 1.04: 0.021; 1.05: 0.022; 1.06: 0.023; 1.07: 0.0215; 1.08: 0.03; 1.10: 0.0225; 1.11: 0.021; 1.12: 0.022; 1.13: 0.022; 1.14: 0.021
| Std Dev. |  | 1.02: 0.001169; 1.03: 0.0005164; 1.04: 0.0013292; 1.05: 0.0019664; 1.06: 0.00075277; 1.07: 0.043765; 1.08: 0; 1.10: 0.00054772; 1.11: 0.0010954; 1.12: 0.0005164; 1.13: 0.0010328; 1.14: 0.00083666
| Rel.Std.Dev. |  | 1.02: 5.12%; 1.03: 2.421%; 1.04: 6.279%; 1.05: 8.675%; 1.06: 3.297%; 1.07: 109.9%; 1.08: 0%; 1.10: 2.434%; 1.11: 5.216%; 1.12: 2.383%; 1.13: 4.767%; 1.14: 3.891%
| PDM3 |  | 1.02: 4.689%; 1.03: -2.188%; 1.04: -2.952%; 1.05: 3.925%; 1.06: 4.689%; 1.07: 82.63%; 1.08: 37.55%; 1.10: 3.161%; 1.11: -3.717%; 1.12: -0.66%; 1.13: -0.66%; 1.14: -1.424%
| Z-Score (Absolute) |  | 1.01: 1.4; 1.02: 0.14; 1.03: 0.72; 1.04: 0.82; 1.05: 0.05; 1.06: 0.14; 1.07: 9.97; 1.08: 4.29; 1.09: 1.4; 1.10: 0.05; 1.11: 0.92; 1.12: 0.53; 1.13: 0.53; 1.14: 0.63
| Analytical results for S in OREAS 505b (Certified Value 0.537 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g |  |  | 3
| 1 | 1 | 0.53 | 0.5 | 0.51 | 0.63 | 0.57 | 0.57 | 0.58 | 0.55 | 0.54 | 0.49077 | 0.57232 | 0.51 | 0.54 | 0.56 | 0.53395 | 0.55 | 0.55 | 0.58 | 0.54 | 0.53 | 0.479 | 0.51785 | 0.51 | 0.56 | 0.51 |  |  | 1
| 1 | 2 | 0.53 | 0.52 | 0.51 | 0.62 | 0.57 | 0.57 | 0.57 | 0.55 | 0.54 | 0.50167 | 0.55829 | 0.5 | 0.53 | 0.57 | 0.55357 | 0.54 | 0.55 | 0.59 | 0.54 | 0.53 | 0.489 | 0.51275 | 0.51 | 0.57 | 0.52 |  |  | 23
| 1 | 3 | 0.52 | 0.53 | 0.51 | 0.63 | 0.57 | 0.57 | 0.58 | 0.55 | 0.54 | 0.4937 | 0.53369 | 0.52 | 0.52 | 0.59 | 0.54172 | 0.53 | 0.55 | 0.58 | 0.54 | 0.54 | 0.493 | 0.51425 | 0.51 | 0.55 | 0.51 |  |  | 16
| 1 | 4 | 0.52 | 0.52 | 0.51 | 0.63 | 0.56 | 0.57 | 0.55 | 0.56 | 0.53 | 0.49083 | 0.56332 | 0.5 | 0.53 | 0.56 | 0.53781 | 0.53 | 0.54 | 0.55 | 0.53 | 0.54 | 0.49 | 0.51425 | 0.51 | 0.56 | 0.52 |  |  | 0.53674
| 1 | 5 | 0.53 | 0.53 | 0.51 | 0.63 | 0.58 | 0.57 | 0.54 | 0.55 | 0.55 | 0.49715 | 0.56756 | 0.51 | 0.53 | 0.58 | 0.55249 | 0.54 | 0.55 | 0.55 | 0.54 | 0.56 | 0.499 | 0.51585 | 0.5 | 0.54 | 0.53 |  |  | 104
| 1 | 6 | 0.53 | 0.52 | 0.51 | 0.63 | 0.56 | 0.56 | 0.53 | 0.56 | 0.54 | 0.48017 | 0.5617 | 0.51 | 0.52 | 0.56 | 0.54507 | 0.54 | 0.53 | 0.56 | 0.56 | 0.58 | 0.494 | 0.51435 | 0.49 | 0.58 | 0.52
| Mean |  | 0.52667 | 0.52 | 0.51 | 0.62833 | 0.56833 | 0.56833 | 0.55833 | 0.55333 | 0.54 | 0.49238 | 0.55948 | 0.50833 | 0.52833 | 0.57 | 0.5441 | 0.53833 | 0.545 | 0.56833 | 0.54167 | 0.54667 | 0.49067 | 0.51488 | 0.505 | 0.56 | 0.51833
| Median |  | 0.53 | 0.52 | 0.51 | 0.63 | 0.57 | 0.57 | 0.56 | 0.55 | 0.54 | 0.49226 | 0.56251 | 0.51 | 0.53 | 0.565 | 0.5434 | 0.54 | 0.55 | 0.57 | 0.54 | 0.54 | 0.4915 | 0.5143 | 0.51 | 0.56 | 0.52
| Std Dev. |  | 0.005164 | 0.010954 | 1.2162e-16 | 0.0040825 | 0.0075277 | 0.0040825 | 0.02137 | 0.005164 | 0.0063246 | 0.0072787 | 0.013539 | 0.0075277 | 0.0075277 | 0.012649 | 0.0078654 | 0.0075277 | 0.0083666 | 0.017224 | 0.0098319 | 0.019664 | 0.0067132 | 0.0017514 | 0.0083666 | 0.014142 | 0.0075277
| Rel.Std.Dev. |  | 0.9805% | 2.107% | 2.385e-14% | 0.6497% | 1.325% | 0.7183% | 3.827% | 0.9332% | 1.171% | 1.478% | 2.42% | 1.481% | 1.425% | 2.219% | 1.446% | 1.398% | 1.535% | 3.031% | 1.815% | 3.597% | 1.368% | 0.3402% | 1.657% | 2.525% | 1.452%
| PDM3 |  | -1.876% | -3.118% | -4.981% | 17.07% | 5.887% | 5.887% | 4.024% | 3.092% | 0.6081% | -8.264% | 4.237% | -5.292% | -1.566% | 6.197% | 1.372% | 0.2975% | 1.54% | 5.887% | 0.9186% | 1.85% | -8.583% | -4.071% | -5.913% | 4.334% | -3.429%
| Z-Score (Absolute) |  | 0.47 | 0.67 | 0.99 | 2.7 | 0.83 | 0.83 | 0.52 | 0.36 | 0.05 | 1.53 | 0.55 | 1.04 | 0.41 | 0.88 | 0.08 | 0.1 | 0.1 | 0.83 | 0 | 0.16 | 1.59 | 0.83 | 1.14 | 0.57 | 0.73
| Analytical results for Sb in OREAS 505b (Certified Value 7.61 ppm) | [28 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g |  | 2
| 1 | 1 | 7.01 | 7 | 7.45 | 8.07 | 6.79 | 8.47 | 8.51 | 8.54 | 8.6 | 4.8275 | 6.6938 | 8.42 | 4.67 | 7.5 | 8.38 | 7.2436 | 7.47 | 7.95 | 9.27 | 7.2 | 7.48 | < 5 | 9.804 | 7.54 | 6.04 | 8 |  | 1
| 1 | 2 | 7.37 | 7.2 | 7.23 | 7.81 | 6.87 | 8.33 | 8.35 | 8.41 | 8.62 | 4.7018 | 6.8261 | 8.22 | 7.55 | 7 | 8.1 | 7.3846 | 7.02 | 8.2 | 9.63 | 7.2 | 7.71 | < 5 | 9.62 | 7.44 | 6.26 | 6 |  | 39
| 1 | 3 | 6.98 | 7.7 | 7.22 | 7.92 | 7.06 | 8.47 | 8.55 | 8.51 | 8.52 | 4.6163 | 6.5838 | 8.14 | 6 | 7.1 | 8.36 | 7.7552 | 7.43 | 7.88 | 9.38 | 7.1 | 7.52 | < 5 | 9.7 | 7.59 | 6.06 | 8 |  | 16
| 1 | 4 | 7.17 | 8 | 7.16 | 7.83 | 6.82 | 8.3 | 8.69 | 8.39 | 8.55 | 5.6809 | 6.7599 | 8.43 | 6.23 | 7.2 | 7.68 | 7.2904 | 7.26 | 8.06 | 9.08 | 7.2 | 7.42 | < 5 | 9.823 | 7.58 | 6.17 | 9 |  | 7.6143
| 1 | 5 | 7.31 | 7.4 | 7.16 | 7.52 | 7.02 | 8.23 | 8.73 | 8.19 | 8.59 | 3.2661 | 6.7582 | 8.36 | 5.55 | 6.8 | 8.11 | 7.7353 | 7.18 | 8.01 | 8.93 | 7.4 | 7.82 | < 5 | 9.927 | 7.47 | 6.18 | 9 |  | 105
| 1 | 6 | 7.46 | 7.8 | 7.28 | 7.46 | 7.25 | 8.4 | 8.3 | 7.62 | 8.42 | 5.6337 | 6.6503 | 8.24 | 4.84 | 7 | 7.95 | 7.2494 | 7.36 | 8.03 | 9.13 | 7.2 | 7.54 | < 5 | 9.971 | 7.49 | 6.14 | 9
| Mean |  | 7.2167 | 7.5167 | 7.25 | 7.7683 | 6.9683 | 8.3667 | 8.5217 | 8.2767 | 8.55 | 4.7877 | 6.712 | 8.3017 | 5.8067 | 7.1 | 8.0967 | 7.4431 | 7.2867 | 8.0217 | 9.2367 | 7.2167 | 7.5817 |  | 9.8075 | 7.5183 | 6.1417 | 8.1667
| Median |  | 7.24 | 7.55 | 7.225 | 7.82 | 6.945 | 8.365 | 8.53 | 8.4 | 8.57 | 4.7646 | 6.726 | 8.3 | 5.775 | 7.05 | 8.105 | 7.3375 | 7.31 | 8.02 | 9.2 | 7.2 | 7.53 |  | 9.8135 | 7.515 | 6.155 | 8.5
| Std Dev. |  | 0.19613 | 0.38166 | 0.10807 | 0.23507 | 0.17543 | 0.096885 | 0.17394 | 0.34442 | 0.073212 | 0.87945 | 0.087199 | 0.11873 | 1.0534 | 0.23664 | 0.26265 | 0.23952 | 0.16896 | 0.10834 | 0.24752 | 0.098319 | 0.15184 |  | 0.13265 | 0.061128 | 0.081588 | 1.169
| Rel.Std.Dev. |  | 2.718% | 5.078% | 1.491% | 3.026% | 2.518% | 1.158% | 2.041% | 4.161% | 0.8563% | 18.37% | 1.299% | 1.43% | 18.14% | 3.333% | 3.244% | 3.218% | 2.319% | 1.351% | 2.68% | 1.362% | 2.003% |  | 1.353% | 0.8131% | 1.328% | 14.31%
| PDM3 |  | -5.223% | -1.283% | -4.785% | 2.022% | -8.484% | 9.88% | 11.92% | 8.698% | 12.29% | -37.12% | -11.85% | 9.027% | -23.74% | -6.755% | 6.334% | -2.249% | -4.303% | 5.349% | 21.31% | -5.223% | -0.4292% |  | 28.8% | -1.261% | -19.34% | 7.254%
| Z-Score (Absolute) |  | 0.35 | 0 | 0.31 | 0.29 | 0.64 | 0.99 | 1.17 | 0.88 | 1.2 | 3.17 | 0.94 | 0.91 | 1.99 | 0.48 | 0.67 | 0.09 | 0.27 | 0.59 | 2 | 0.35 | 0.08 | 5.83 | 2.66 | 0 | 1.6 | NA
Results from laboratory 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Sc in OREAS 505b (Certified Value 7.48 ppm) | [28 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.27 | 1.29 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g |  |  |  |  | 3
| 1 | 1 | 7 | 6.6 | 7.3 | 8.3 | 7.7 | 7.5 | 8.8 | 7.9 | 7.65 | 6.4523 | 7.4195 | 7.5 | 7.2 | 7.4 | 9.5 | 6.8579 | 7.3 | 7.8 | 8.5 | 7.2 | 6.8 | 7.9 | 7.2 |  |  |  |  | 1
| 1 | 2 | 7 | 6.8 | 7.1 | 8 | 7.7 | 7.7 | 8 | 7.6 | 7.65 | 6.2982 | 7.2234 | 7.5 | 7.15 | 7.6 | 9.6 | 6.8292 | 6.8 | 7.9 | 8.8 | 7.3 | 6.7 | 8 | 7.2 |  |  |  |  | 40
| 1 | 3 | 7 | 6.9 | 7.3 | 8.2 | 8.1 | 7.9 | 8.1 | 7.7 | 7.84 | 6.3105 | 7.0957 | 7.7 | 7.2 | 7.5 | 9.5 | 6.8328 | 7.3 | 7.9 | 8.3 | 7.3 | 6.7 | 8 | 6.4 |  |  |  |  | 16
| 1 | 4 | 7 | 7 | 7.1 | 8.2 | 7.9 | 7.7 | 8.5 | 7.7 | 7.72 | 6.1755 | 7.274 | 7.7 | 7.05 | 7.4 | 9.2 | 6.8347 | 7.2 | 8.1 | 8.4 | 7.3 | 6.9 | 8.1 | 7.9 |  |  |  |  | 7.4821
| 1 | 5 | 7 | 7 | 7 | 8.1 | 7.9 | 7.5 | 8.4 | 7.7 | 8.07 | 6.9007 | 7.114 | 7.4 | 7.56 | 7.3 | 9.6 | 6.9891 | 7.3 | 8.2 | 8.2 | 7.3 | 6.7 | 7.9 | 6.5 |  |  |  |  | 106
| 1 | 6 | 7 | 6.9 | 7.5 | 7.7 | 8 | 7.4 | 8.2 | 7.3 | 7.87 | 6.494 | 7.1628 | 7.7 | 6.97 | 7.5 | 9.2 | 7.0072 | 7.5 | 8 | 8.4 | 7.4 | 6.8 | 8.1 | 7
| Mean |  | 7 | 6.8667 | 7.2167 | 8.0833 | 7.8833 | 7.6167 | 8.3333 | 7.65 | 7.8 | 6.4385 | 7.2149 | 7.5833 | 7.1883 | 7.45 | 9.4333 | 6.8918 | 7.2333 | 7.9833 | 8.4333 | 7.3 | 6.7667 | 8 | 7.0333
| Median |  | 7 | 6.9 | 7.2 | 8.15 | 7.9 | 7.6 | 8.3 | 7.7 | 7.78 | 6.3814 | 7.1931 | 7.6 | 7.175 | 7.45 | 9.5 | 6.8463 | 7.3 | 7.95 | 8.4 | 7.3 | 6.75 | 8 | 7.1
| Std Dev. |  | 0 | 0.15055 | 0.18348 | 0.2137 | 0.16021 | 0.18348 | 0.29439 | 0.19748 | 0.16174 | 0.25383 | 0.12044 | 0.13292 | 0.20331 | 0.10488 | 0.18619 | 0.083184 | 0.23381 | 0.1472 | 0.20656 | 0.063246 | 0.08165 | 0.089443 | 0.5465
| Rel.Std.Dev. |  | 0% | 2.193% | 2.543% | 2.644% | 2.032% | 2.409% | 3.533% | 2.581% | 2.074% | 3.942% | 1.669% | 1.753% | 2.828% | 1.408% | 1.974% | 1.207% | 3.232% | 1.844% | 2.449% | 0.8664% | 1.207% | 1.118% | 7.77%
| PDM3 |  | -6.443% | -8.225% | -3.547% | 8.036% | 5.363% | 1.799% | 11.38% | 2.244% | 4.249% | -13.95% | -3.571% | 1.353% | -3.926% | -0.4287% | 26.08% | -7.889% | -3.324% | 6.699% | 12.71% | -2.433% | -9.562% | 6.922% | -5.998%
| Z-Score (Absolute) |  | NA | 1.05 | 0.49 | 0.92 | 0.59 | 0.16 | 1.32 | 0.22 | 0.46 | 1.74 | 0.49 | 0.11 | 0.53 | 0.11 | 3.1 | 1.01 | 0.46 | 0.76 | 1.48 | 0.35 | 1.21 | 0.78 | 0.78
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 505b (Certified Value 3.72 ppm) | [28 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 01g | 0.5g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | [5 empty] | 3
| 1 | 1 | 3 | 3.5 | 3.7 | 4 | 3.9 | 3.8 | 3.5 | 3.5 | 3.9038 | 90.7 | 3.54 | 3 | 3.4617 | 3.9 | 3.8 | 3.7 | < 5 | 3.8 | < 10 | 3.52 | 4.4 | 4 | [5 empty] | 1
| 1 | 2 | 3 | 3.4 | 3.7 | 3.8 | 3.8 | 3.7 | 3.7 | 3.55 | 3.7711 | 5.6 | 2.54 | 3 | 4.2586 | 3.5 | 4.3 | 3.5 | < 5 | 4.1 | < 10 | 3.597 | 4.6 | 4 | [5 empty] | 41
| 1 | 3 | 4 | 3.8 | 3.7 | 4.3 | 3.9 | 3.7 | 4 | 3.47 | 3.8492 | 6.7 | 0.92 | 3 | 3.6295 | 3.6 | 3.7 | 3.3 | < 5 | 3.4 | < 10 | 3.524 | 4.2 | 3 | [5 empty] | 16
| 1 | 4 | 4 | 3.3 | 3.7 | 3.7 | 3.9 | 3.8 | 3.6 | 3.45 | 3.7045 | 1.7 | 1.86 | 2 | 3.6932 | 3.6 | 3.7 | 3.5 | < 5 | 3.3 | < 10 | 3.582 | 4.5 | 4 | [5 empty] | 3.7153
| 1 | 5 | 3 | 3.5 | 3.7 | 4.2 | 4.2 | 3.6 | 4.2 | 3.72 | 3.8312 | 6 | 4.24 | 3 | 3.7821 | 3.6 | 3.6 | 3.6 | < 5 | 3.2 | < 10 | 3.737 | 5 | 4 | [5 empty] | 107
| 1 | 6 | 4 | 3.8 | 3.6 | 4.2 | 4.2 | 4 | 3.4 | 3.48 | 4.0706 | 4.6 | 2 | 3 | 4.248 | 3.7 | 3.9 | 3.7 | < 5 | 3.7 | < 10 | 3.638 | 4.6 | 3
| Mean |  | 3.5 | 3.55 | 3.6833 | 4.0333 | 3.9833 | 3.7667 | 3.7333 | 3.5283 | 3.8551 | 19.2167 | 2.5167 | 2.8333 | 3.8455 | 3.65 | 3.8333 | 3.55 |  | 3.5833 |  | 3.5997 | 4.55 | 3.6667
| Median |  | 3.5 | 3.5 | 3.7 | 4.1 | 3.9 | 3.75 | 3.65 | 3.49 | 3.8402 | 5.8 | 2.27 | 3 | 3.7376 | 3.6 | 3.75 | 3.55 |  | 3.55 |  | 3.5895 | 4.55 | 4
| Std Dev. |  | 0.54772 | 0.20736 | 0.040825 | 0.24221 | 0.17224 | 0.13663 | 0.30768 | 0.099883 | 0.12583 | 35.0631 | 1.2054 | 0.40825 | 0.3328 | 0.13784 | 0.25033 | 0.15166 |  | 0.34303 |  | 0.080897 | 0.26646 | 0.5164
| Rel.Std.Dev. |  | 15.65% | 5.841% | 1.108% | 6.005% | 4.324% | 3.627% | 8.241% | 2.831% | 3.264% | 182.5% | 47.9% | 14.41% | 8.654% | 3.776% | 6.53% | 4.272% |  | 9.573% |  | 2.247% | 5.856% | 14.08%
| PDM3 |  | -5.794% | -4.448% | -0.8592% | 8.561% | 7.216% | 1.384% | 0.4866% | -5.031% | 3.763% | 417.2% | -32.26% | -23.74% | 3.506% | -1.756% | 3.178% | -4.448% |  | -3.551% |  | -3.111% | 22.47% | -1.308%
| Z-Score (Absolute) |  | NA | 0.67 | 0.18 | 1.1 | 0.92 | 0.12 | 0 | 0.75 | 0.45 | 56.95 | 4.47 | NA | 0.41 | 0.31 | 0.37 | 0.67 | 4.54 | 0.55 | 4.66 | 0.49 | 3 | NA
Results from laboratories 1.01, 1.17 and 1.29 were removed due to their 1 ppm reading resolution.
| Analytical results for Si in OREAS 505b (Indicative Value 0.029 wt.%) | [28 empty] | Indicative
Si | Round | Replicate | 1.01: Lab | [5 empty] | 1
| No | No | 1.01: 1.13 | [5 empty] | wt.%
|  |  | 1.01: AR*OES/MS | [5 empty] | 3
|  |  | 1.01: 0.5g | [5 empty] | 3
| 1 | 1 | 1.01: 0.025482 | [5 empty] | 1
| 1 | 2 | 1.01: 0.037531 | [5 empty] | 4
| 1 | 3 | 1.01: 0.024365 | [5 empty] | 16
| 1 | 4 | 1.01: 0.023289 | [5 empty] | 0.028597
| 1 | 5 | 1.01: 0.024928 | [5 empty] | 12
| 1 | 6 | 1.01: 0.035985
| Mean |  | 1.01: 0.028597
| Median |  | 1.01: 0.025205
| Std Dev. |  | 1.01: 0.006382
| Rel.Std.Dev. |  | 1.01: 22.32%
| PDM3 |  | 1.01: -1.11e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Sm in OREAS 505b (Indicative Value 3.38 ppm) | [28 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | [5 empty] | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 | [5 empty] | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS | [5 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g | [5 empty] | 2
| 1 | 1 | 1.01: 3.393; 1.02: 3.1798; 1.03: 4.3131 | [5 empty] | 1
| 1 | 2 | 1.01: 3.312; 1.02: 3.2417; 1.03: 4.2658 | [5 empty] | 7
| 1 | 3 | 1.01: 3.351; 1.02: 3.318; 1.03: 4.2744 | [5 empty] | 16
| 1 | 4 | 1.01: 3.232; 1.02: 3.3879; 1.03: 4.2223 | [5 empty] | 3.3838
| 1 | 5 | 1.01: 3.394; 1.02: 3.705; 1.03: 4.2156 | [5 empty] | 13
| 1 | 6 | 1.01: 3.464; 1.02: 3.6275; 1.03: 4.2287
| Mean |  | 1.01: 3.3577; 1.02: 3.41; 1.03: 4.2533
| Median |  | 1.01: 3.372; 1.02: 3.3529; 1.03: 4.2472
| Std Dev. |  | 1.01: 0.079746; 1.02: 0.21194; 1.03: 0.037846
| Rel.Std.Dev. |  | 1.01: 2.375%; 1.02: 6.215%; 1.03: 0.8898%
| PDM3 |  | 1.01: -0.773%; 1.02: 0.773%; 1.03: 25.7%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0; 1.03: 10.87
| Analytical results for Sn in OREAS 505b (Certified Value 3.13 ppm) | [28 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.30 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | [5 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | [5 empty] | 3
| 1 | 1 | 3.1 | 3.3 | 2.8 | 3.3 | 3.4 | 3.3 | 3.19 | 3.945 | 3.263 | 7.4 | 3.6 | < 5 | 4 | 2.6793 | 2.8 | 2.9 | 3.3 | < 10 | < 10 | 3.1 | 3.1 | < 10 | [5 empty] | 1
| 1 | 2 | 3.1 | 3.2 | 2.8 | 3.3 | 3.2 | 3.1 | 3.15 | 3.9493 | 3.2368 | 3.3 | 5.75 | < 5 | 3.8 | 2.9501 | 2.6 | 3 | 3.3 | < 10 | < 10 | 3.1 | 3.4 | < 10 | [5 empty] | 14
| 1 | 3 | 3.2 | 3.3 | 2.8 | 3.3 | 3.3 | 3.2 | 3.14 | 3.5879 | 3.2307 | 3.2 | 3.59 | < 5 | 3.8 | 2.695 | 2.8 | 2.9 | 3.2 | < 10 | < 10 | 3.1 | 3.2 | < 10 | [5 empty] | 16
| 1 | 4 | 3.1 | 3.3 | 2.8 | 3.3 | 3.6 | 3.3 | 3.15 | 3.9076 | 3.2707 | 3.1 | 4.91 | < 5 | 3.7 | 2.8238 | 2.7 | 3 | 3.1 | < 10 | < 10 | 3.1 | 3.4 | < 10 | [5 empty] | 3.1256
| 1 | 5 | 3.1 | 3.2 | 2.8 | 3.3 | 3.4 | 3.3 | 3.13 | 3.5982 | 3.2147 | 3.7 | 5.31 | < 5 | 3.8 | 2.961 | 2.8 | 3 | 3.1 | < 10 | < 10 | 3 | 3.3 | < 10 | [5 empty] | 108
| 1 | 6 | 3.4 | 3.2 | 2.8 | 3.1 | 3.4 | 3 | 3.16 | 3.5273 | 3.2305 | 3.4 | 5.3 | < 5 | 3.7 | 2.8662 | 2.8 | 3 | 3.2 | < 10 | < 10 | 3.1 | 3.4 | < 10
| Mean |  | 3.1667 | 3.25 | 2.8 | 3.2667 | 3.3833 | 3.2 | 3.1533 | 3.7526 | 3.2411 | 4.0167 | 4.7433 |  | 3.8 | 2.8292 | 2.75 | 2.9667 | 3.2 |  |  | 3.0833 | 3.3
| Median |  | 3.1 | 3.25 | 2.8 | 3.3 | 3.4 | 3.25 | 3.15 | 3.7529 | 3.2338 | 3.35 | 5.105 |  | 3.8 | 2.845 | 2.8 | 3 | 3.2 |  |  | 3.1 | 3.35
| Std Dev. |  | 0.12111 | 0.054772 | 4.8648e-16 | 0.08165 | 0.13292 | 0.12649 | 0.020656 | 0.20075 | 0.021412 | 1.6702 | 0.92839 |  | 0.10954 | 0.12154 | 0.083666 | 0.05164 | 0.089443 |  |  | 0.040825 | 0.12649
| Rel.Std.Dev. |  | 3.824% | 1.685% | 1.737e-14% | 2.499% | 3.929% | 3.953% | 0.6551% | 5.35% | 0.6606% | 41.58% | 19.57% |  | 2.883% | 4.296% | 3.042% | 1.741% | 2.795% |  |  | 1.324% | 3.833%
| PDM3 |  | 1.315% | 3.981% | -10.42% | 4.514% | 8.247% | 2.381% | 0.8881% | 20.06% | 3.695% | 28.51% | 51.76% |  | 21.58% | -9.481% | -12.02% | -5.084% | 2.381% |  |  | -1.352% | 5.581%
| Z-Score (Absolute) |  | 0.15 | 0.01 | 0.86 | 0.04 | 0.27 | 0.09 | 0.18 | 0.98 | 0.01 | 1.5 | 2.91 | 1.45 | 1.08 | 0.81 | 0.96 | 0.54 | 0.09 | 3.4 | 3.4 | 0.31 | 0.11 | 3.4
| Analytical results for Sr in OREAS 505b (Certified Value 91 ppm) | [28 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 0
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 1
| 1 | 1 | 90 | 82 | 86.1 | 94.1 | 94.9 | 91.7 | 99 | 93.3 | 92.88 | 94.9242 | 90.6999 | 86.1 | 87.39 | 93.5 | 86.4 | 97.8 | 87.8367 | 88.5 | 91.5 | 100 | 83.9 | 93.5 | 83.93 | 91 | 94.2 | 85.8 |  | 1
| 1 | 2 | 91 | 83.5 | 84.7 | 93.7 | 97.1 | 91.1 | 94.8 | 91.1 | 92.11 | 91.4801 | 89.0608 | 85.8 | 83.64 | 90.21 | 88.6 | 98.7 | 89.9954 | 87.7 | 93.4 | 103 | 83 | 94.1 | 84.67 | 90 | 98.3 | 86.3 |  | 15
| 1 | 3 | 90.9 | 87.6 | 84.7 | 94.8 | 98.4 | 92.7 | 98.1 | 92.6 | 91.46 | 92.0438 | 88.8697 | 86.5 | 83.82 | 93.39 | 88.5 | 99.3 | 86.8327 | 92.9 | 92.9 | 100 | 84.1 | 93.7 | 85.42 | 92 | 94 | 85.1 |  | 16
| 1 | 4 | 89.4 | 85.8 | 84.8 | 95.2 | 93.7 | 90.8 | 92.5 | 91.7 | 90.19 | 91.3432 | 91.4098 | 85.6 | 85.9 | 93.82 | 87.2 | 98.4 | 86.9784 | 91.9 | 91.2 | 95.8 | 83.2 | 91.7 | 85.92 | 91 | 97 | 84.8 |  | 90.7258
| 1 | 5 | 90.6 | 86.5 | 85.1 | 95.4 | 99.4 | 90.4 | 87.8 | 93.3 | 92.69 | 93.543 | 89.7678 | 87.5 | 83.19 | 92.24 | 86.4 | 98.2 | 93.9164 | 92.7 | 95.9 | 95.4 | 82.1 | 95.5 | 85.57 | 88 | 95.7 | 87 |  | 109
| 1 | 6 | 91.8 | 84.5 | 85.6 | 94.5 | 97.4 | 88.6 | 88.4 | 89 | 92.57 | 92.0012 | 89.6455 | 85.6 | 82.57 | 93.58 | 88.3 | 97.3 | 93.9505 | 90.4 | 93.8 | 97.4 | 83.3 | 92.3 | 87.71 | 88 | 98.7 | 86.3
| Mean |  | 90.6167 | 84.9833 | 85.1667 | 94.6167 | 96.8167 | 90.8833 | 93.4333 | 91.8333 | 91.9833 | 92.5559 | 89.9089 | 86.1833 | 84.4183 | 92.79 | 87.5667 | 98.2833 | 89.9183 | 90.6833 | 93.1167 | 98.6 | 83.2667 | 93.4667 | 85.5367 | 90 | 96.3167 | 85.8833
| Median |  | 90.75 | 85.15 | 84.95 | 94.65 | 97.25 | 90.95 | 93.65 | 92.15 | 92.34 | 92.0225 | 89.7066 | 85.95 | 83.73 | 93.445 | 87.75 | 98.3 | 88.916 | 91.15 | 93.15 | 98.7 | 83.25 | 93.6 | 85.495 | 90.5 | 96.35 | 86.05
| Std Dev. |  | 0.83526 | 2.0566 | 0.57155 | 0.64936 | 2.1442 | 1.3761 | 4.745 | 1.6415 | 1.0152 | 1.3987 | 0.97641 | 0.73052 | 1.8393 | 1.379 | 1.0328 | 0.6969 | 3.309 | 2.2004 | 1.7105 | 2.9257 | 0.71181 | 1.3471 | 1.282 | 1.6733 | 2.0154 | 0.82321
| Rel.Std.Dev. |  | 0.9218% | 2.42% | 0.6711% | 0.6863% | 2.215% | 1.514% | 5.078% | 1.788% | 1.104% | 1.511% | 1.086% | 0.8476% | 2.179% | 1.486% | 1.179% | 0.7091% | 3.68% | 2.426% | 1.837% | 2.967% | 0.8549% | 1.441% | 1.499% | 1.859% | 2.092% | 0.9585%
| PDM3 |  | -0.1203% | -6.329% | -6.127% | 4.289% | 6.714% | 0.1736% | 2.984% | 1.221% | 1.386% | 2.017% | -0.9004% | -5.007% | -6.952% | 2.275% | -3.482% | 8.33% | -0.89% | -0.0468% | 2.635% | 8.679% | -8.222% | 3.021% | -5.72% | -0.8% | 6.162% | -5.337%
| Z-Score (Absolute) |  | 0.04 | 1.33 | 1.28 | 0.88 | 1.38 | 0.02 | 0.61 | 0.24 | 0.27 | 0.41 | 0.2 | 1.05 | 1.45 | 0.46 | 0.74 | 1.71 | 0.2 | 0.02 | 0.53 | 1.79 | 1.72 | 0.61 | 1.2 | 0.18 | 1.26 | 1.12
| Analytical results for Ta in OREAS 505b (Certified Value < 0.01 ppm) | [28 empty] | Certified
Ta | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.16: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.05; 1.03: 1.06; 1.04: 1.07; 1.05: 1.08; 1.06: 1.09; 1.07: 1.12; 1.08: 1.17; 1.09: 1.18; 1.10: 1.19; 1.11: 1.20; 1.12: 1.21; 1.13: 1.22; 1.14: 1.27; 1.16: 1.29 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*MS; 1.12: AR*MS; 1.13: AR*OES; 1.14: AR*OES/MS; 1.16: AR*OES/MS |  | 3
|  |  | 1.01: 01g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.25g; 1.09: 01g; 1.10: 0.2g; 1.11: 0.5g; 1.12: 0.5g; 1.13: 0.5g; 1.14: 0.25g; 1.16: 0.25g |  | 3
| 1 | 1 | 1.01: 0.56; 1.02: 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.06: < 0.01; 1.07: 0.11; 1.08: < 0.05; 1.09: < 0.05; 1.10: < 0.01; 1.11: 0.01; 1.12: 0.01; 1.13: < 5; 1.14: < 0.1; 1.16: < 0.05 |  | 1
| 1 | 2 | 1.01: 0.51; 1.02: < 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: < 0.01; 1.06: < 0.01; 1.07: 0.07; 1.08: < 0.05; 1.09: < 0.05; 1.10: < 0.01; 1.11: 0.01; 1.12: 0.01; 1.13: < 5; 1.14: < 0.1; 1.16: < 0.05 |  | 28
| 1 | 3 | 1.01: 0.56; 1.02: < 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.06: < 0.01; 1.07: 0.09; 1.08: < 0.05; 1.09: < 0.05; 1.10: < 0.01; 1.11: < 0.01; 1.12: 0.01; 1.13: < 5; 1.14: < 0.1; 1.16: < 0.05 |  | 16
| 1 | 4 | 1.01: 0.6; 1.02: 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: < 0.01; 1.06: < 0.01; 1.07: 0.07; 1.08: < 0.05; 1.09: < 0.05; 1.10: < 0.01; 1.11: 0.01; 1.12: 0.01; 1.13: < 5; 1.14: < 0.1; 1.16: < 0.05 |  | < 0.01
| 1 | 5 | 1.01: 0.56; 1.02: < 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.06: < 0.01; 1.07: 0.06; 1.08: < 0.05; 1.09: < 0.05; 1.10: < 0.01; 1.11: 0.01; 1.12: 0.06; 1.13: < 5; 1.14: < 0.1; 1.16: < 0.05 |  | 110
| 1 | 6 | 1.01: 0.57; 1.02: < 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: < 0.01; 1.06: < 0.01; 1.07: 0.05; 1.08: < 0.05; 1.09: < 0.05; 1.10: < 0.01; 1.11: 0.01; 1.12: 0.01; 1.13: < 5; 1.14: < 0.1; 1.16: < 0.05
| Mean |  | 1.01: 0.56; 1.02: 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.07: 0.075; 1.11: 0.01; 1.12: 0.018333
| Median |  | 1.01: 0.56; 1.02: 0.01; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.07: 0.07; 1.11: 0.01; 1.12: 0.01
| Std Dev. |  | 1.01: 0.028983; 1.02: 0; 1.03: 0; 1.04: 0; 1.05: 0; 1.07: 0.021679; 1.11: 0; 1.12: 0.020412
| Rel.Std.Dev. |  | 1.01: 5.175%; 1.02: 0%; 1.03: 0%; 1.04: 0%; 1.05: 0%; 1.07: 28.91%; 1.11: 0%; 1.12: 111.3%
PDM3
| Z-Score (Absolute) |  | 1.01: 33.72; 1.02: 0.73; 1.03: 0.52; 1.04: 0.52; 1.05: 0.67; 1.06: 0.83; 1.07: 3.53; 1.08: 0.41; 1.09: 0.41; 1.10: 0.83; 1.11: 0.57; 1.12: 0; 1.13: 154.47; 1.14: 1.97; 1.16: 0.41
| Analytical results for Tb in OREAS 505b (Certified Value 0.5 ppm) | [28 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11; 1.04: 1.17; 1.05: 1.29 |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*OES/MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g; 1.04: 0.25g; 1.05: 0.25g |  | 2
| 1 | 1 | 1.01: 0.477; 1.02: 0.4128; 1.03: 0.51978; 1.04: 0.6; 1.05: 0.43 |  | 1
| 1 | 2 | 1.01: 0.477; 1.02: 0.4268; 1.03: 0.51396; 1.04: 0.62; 1.05: 0.45 |  | 29
| 1 | 3 | 1.01: 0.475; 1.02: 0.4288; 1.03: 0.50744; 1.04: 0.63; 1.05: 0.43 |  | 16
| 1 | 4 | 1.01: 0.475; 1.02: 0.4354; 1.03: 0.5162; 1.04: 0.59; 1.05: 0.45 |  | 0.49519
| 1 | 5 | 1.01: 0.461; 1.02: 0.4743; 1.03: 0.50418; 1.04: 0.61; 1.05: 0.46 |  | 111
| 1 | 6 | 1.01: 0.464; 1.02: 0.4579; 1.03: 0.50902; 1.04: 0.61; 1.05: 0.44
| Mean |  | 1.01: 0.4715; 1.02: 0.43933; 1.03: 0.51176; 1.04: 0.61; 1.05: 0.44333
| Median |  | 1.01: 0.475; 1.02: 0.4321; 1.03: 0.51149; 1.04: 0.61; 1.05: 0.445
| Std Dev. |  | 1.01: 0.0070922; 1.02: 0.022605; 1.03: 0.0058724; 1.04: 0.014142; 1.05: 0.012111
| Rel.Std.Dev. |  | 1.01: 1.504%; 1.02: 5.145%; 1.03: 1.147%; 1.04: 2.318%; 1.05: 2.732%
| PDM3 |  | 1.01: -4.783%; 1.02: -11.28%; 1.03: 3.348%; 1.04: 23.19%; 1.05: -10.47%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.03: 0.84; 1.04: 2.9; 1.05: 0.59
| Analytical results for Te in OREAS 505b (Certified Value 0.16 ppm) | [28 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 15g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | [6 empty] | 2
| 1 | 1 | 0.15 | < 0.2 | 0.17 | 0.16 | 0.17 | 0.16 | 0.15 | 0.14 | < 5 | 46.2 | 0.17 | 0.2014 | 0.18 | 0.17 | 0.16 | 6.4 | 0.22 | < 10 | 0.244 | 0.2 | 0.19 | [6 empty] | 1
| 1 | 2 | 0.1 | < 0.2 | 0.17 | 0.16 | 0.16 | 0.17 | 0.19 | 0.14 | < 5 | 2.54 | 0.15 | 0.234 | 0.17 | 0.16 | 0.18 | 6.5 | 0.07 | < 10 | 0.239 | 0.2 | 0.16 | [6 empty] | 30
| 1 | 3 | 0.09 | < 0.2 | 0.17 | 0.16 | 0.17 | 0.18 | 0.16 | 0.14 | < 5 | 2.22 | 0.17 | 0.21273 | 0.18 | 0.15 | 0.15 | 6.5 | 0.2 | < 10 | 0.2414 | 0.1 | 0.16 | [6 empty] | 16
| 1 | 4 | 0.09 | < 0.2 | 0.18 | 0.16 | 0.17 | 0.15 | 0.12 | 0.14 | < 5 | 2.3 | 0.15 | 0.23386 | 0.17 | 0.16 | 0.16 | 6.6 | 0.11 | < 10 | 0.252 | 0.2 | 0.15 | [6 empty] | 0.16436
| 1 | 5 | 0.11 | < 0.2 | 0.17 | 0.17 | 0.18 | 0.17 | 0.23 | 0.15 | < 5 | 3.13 | 0.15 | 0.20208 | 0.16 | 0.18 | 0.15 | 6.6 | 0.15 | < 10 | 0.251 | 0.2 | 0.17 | [6 empty] | 112
| 1 | 6 | 0.13 | < 0.2 | 0.19 | 0.15 | 0.16 | 0.15 | 0.21 | 0.14 | < 5 | 1.98 | 0.13 | 0.23562 | 0.19 | 0.17 | 0.14 | 6.7 | 0.19 | < 10 | 0.248 | 0.1 | 0.15
| Mean |  | 0.11167 |  | 0.175 | 0.16 | 0.16833 | 0.16333 | 0.17667 | 0.14167 |  | 9.7283 | 0.15333 | 0.21995 | 0.175 | 0.165 | 0.15667 | 6.55 | 0.15667 |  | 0.2459 | 0.16667 | 0.16333
| Median |  | 0.105 |  | 0.17 | 0.16 | 0.17 | 0.165 | 0.175 | 0.14 |  | 2.42 | 0.15 | 0.2233 | 0.175 | 0.165 | 0.155 | 6.55 | 0.17 |  | 0.246 | 0.2 | 0.16
| Std Dev. |  | 0.024014 |  | 0.0083666 | 0.0063246 | 0.0075277 | 0.012111 | 0.040825 | 0.0040825 |  | 17.8717 | 0.015055 | 0.016446 | 0.010488 | 0.010488 | 0.013663 | 0.10488 | 0.05785 |  | 0.0052745 | 0.05164 | 0.015055
| Rel.Std.Dev. |  | 21.5% |  | 4.781% | 3.953% | 4.472% | 7.415% | 23.11% | 2.882% |  | 183.7% | 9.819% | 7.477% | 5.993% | 6.356% | 8.721% | 1.601% | 36.93% |  | 2.145% | 30.98% | 9.218%
| PDM3 |  | -32.06% |  | 6.473% | -2.653% | 2.417% | -0.6253% | 7.487% | -13.81% |  | 5819% | -6.709% | 33.82% | 6.473% | 0.3887% | -4.681% | 3885% | -4.681% |  | 49.61% | 1.403% | -0.6253%
| Z-Score (Absolute) |  | 3.08 | 3.76 | 0.58 | 0.29 | 0.19 | 0.1 | 0.67 | 1.35 | 134.96 | 552.74 | 0.67 | 3.18 | 0.58 | 0 | 0.48 | NA | 0.48 | 279.45 | 4.68 | NA | 0.1
Results from laboratories 1.22 and 1.27 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Th in OREAS 505b (Certified Value 6.48 ppm) | [28 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | [6 empty] | 2
| 1 | 1 | 10.5 | 5.5 | 5.2 | 6.8 | 7.2 | 6.2 | 7.6 | 6.5 | 5.955 | 6.6684 | 6.301 | 5.8 | 7.9 | 5.2 | 7.223 | 7.6 | 6.4 | 6.5 | 5 | 6.6 | 6.3 | [6 empty] | 1
| 1 | 2 | 10.6 | 5.9 | 4.8 | 6.9 | 6.8 | 6.3 | 7.3 | 6.4 | 5.856 | 6.9543 | 6.38 | 5.4 | 8.1 | 5.3 | 7.7061 | 6.9 | 6 | 6.6 | 5.2 | 6.4 | 6.5 | [6 empty] | 31
| 1 | 3 | 10.7 | 5.9 | 5.1 | 7.1 | 7 | 6.5 | 7.6 | 6.3 | 5.747 | 7.0839 | 6.233 | 5.6 | 8.5 | 5.6 | 7.6519 | 7.5 | 6.1 | 6.4 | 5.4 | 6.5 | 6.3 | [6 empty] | 16
| 1 | 4 | 10.9 | 6.2 | 5.2 | 7.6 | 6.9 | 6.3 | 7.8 | 6.5 | 5.679 | 7.111 | 6.5372 | 5.7 | 8.3 | 5.2 | 7.6307 | 7.4 | 6.3 | 6 | 5.1 | 6.3 | 6.4 | [6 empty] | 6.4778
| 1 | 5 | 10.9 | 5.9 | 5 | 7.5 | 7 | 6.3 | 7.6 | 6.3 | 5.817 | 7.8216 | 6.2157 | 5.5 | 8.5 | 5.9 | 7.5856 | 7.1 | 6.4 | 6 | 5.4 | 6.4 | 6.6 | [6 empty] | 113
| 1 | 6 | 10.6 | 5.9 | 5.1 | 7.1 | 6.7 | 6.4 | 7.5 | 6.1 | 5.942 | 8.0974 | 6.1237 | 5.5 | 8.2 | 5.5 | 7.2531 | 7.4 | 6.3 | 6.3 | 5.2 | 6.3 | 6.7
| Mean |  | 10.7 | 5.8833 | 5.0667 | 7.1667 | 6.9333 | 6.3333 | 7.5667 | 6.35 | 5.8327 | 7.2894 | 6.2984 | 5.5833 | 8.25 | 5.45 | 7.5084 | 7.3167 | 6.25 | 6.3 | 5.2167 | 6.4167 | 6.4667
| Median |  | 10.65 | 5.9 | 5.1 | 7.1 | 6.95 | 6.3 | 7.6 | 6.35 | 5.8365 | 7.0975 | 6.267 | 5.55 | 8.25 | 5.4 | 7.6081 | 7.4 | 6.3 | 6.35 | 5.2 | 6.4 | 6.45
| Std Dev. |  | 0.16733 | 0.22286 | 0.15055 | 0.32042 | 0.17512 | 0.10328 | 0.1633 | 0.15166 | 0.10832 | 0.5492 | 0.14516 | 0.1472 | 0.23452 | 0.27386 | 0.21316 | 0.26394 | 0.16432 | 0.25298 | 0.16021 | 0.1169 | 0.1633
| Rel.Std.Dev. |  | 1.564% | 3.788% | 2.971% | 4.471% | 2.526% | 1.631% | 2.158% | 2.388% | 1.857% | 7.534% | 2.305% | 2.636% | 2.843% | 5.025% | 2.839% | 3.607% | 2.629% | 4.016% | 3.071% | 1.822% | 2.525%
| PDM3 |  | 65.18% | -9.177% | -21.78% | 10.63% | 7.033% | -2.23% | 16.81% | -1.973% | -9.959% | 12.53% | -2.769% | -13.81% | 27.36% | -15.87% | 15.91% | 12.95% | -3.516% | -2.744% | -19.47% | -0.9434% | -0.1716%
| Z-Score (Absolute) |  | 3.83 | 0.41 | 1.13 | 0.72 | 0.51 | 0.01 | 1.07 | 0 | 0.46 | 0.83 | 0.05 | 0.67 | 1.67 | 0.79 | 1.02 | 0.85 | 0.09 | 0.04 | 1 | 0.06 | 0.1
| Analytical results for Ti in OREAS 505b (Certified Value 0.238 wt.%) | [28 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 3
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  |  | 3
| 1 | 1 | 0.236 | 0.23 | 0.236 | 0.25 | 0.249 | 0.233 | 0.247 | 0.236 | 0.2274 | 0.21058 | 0.22723 | 0.231 | 0.24038 | 0.26 | 0.25 | 0.28458 | 0.22 | 0.241 | 0.257 | 0.24 | 0.217 | 0.2345 | 0.208 | 0.26 | 0.231 |  |  | 1
| 1 | 2 | 0.236 | 0.235 | 0.227 | 0.25 | 0.246 | 0.229 | 0.246 | 0.238 | 0.2276 | 0.20902 | 0.22071 | 0.235 | 0.23089 | 0.27 | 0.26 | 0.29344 | 0.22 | 0.243 | 0.263 | 0.24 | 0.219 | 0.23769 | 0.208 | 0.25 | 0.232 |  |  | 32
| 1 | 3 | 0.237 | 0.242 | 0.232 | 0.25 | 0.243 | 0.228 | 0.245 | 0.24 | 0.2289 | 0.21223 | 0.21995 | 0.233 | 0.23619 | 0.27 | 0.26 | 0.28668 | 0.22 | 0.241 | 0.252 | 0.24 | 0.235 | 0.23341 | 0.212 | 0.26 | 0.229 |  |  | 16
| 1 | 4 | 0.238 | 0.228 | 0.231 | 0.25 | 0.246 | 0.228 | 0.242 | 0.238 | 0.2216 | 0.2095 | 0.22397 | 0.231 | 0.23656 | 0.26 | 0.25 | 0.28502 | 0.22 | 0.241 | 0.246 | 0.24 | 0.227 | 0.23263 | 0.206 | 0.26 | 0.256 |  |  | 0.23753
| 1 | 5 | 0.238 | 0.228 | 0.231 | 0.25 | 0.245 | 0.229 | 0.24 | 0.239 | 0.2255 | 0.2086 | 0.22114 | 0.236 | 0.2317 | 0.26 | 0.26 | 0.29206 | 0.22 | 0.245 | 0.243 | 0.24 | 0.237 | 0.24059 | 0.202 | 0.25 | 0.23 |  |  | 114
| 1 | 6 | 0.243 | 0.238 | 0.234 | 0.25 | 0.243 | 0.231 | 0.238 | 0.243 | 0.2306 | 0.20988 | 0.22122 | 0.23 | 0.23117 | 0.26 | 0.25 | 0.28643 | 0.22 | 0.237 | 0.254 | 0.25 | 0.237 | 0.24117 | 0.203 | 0.26 | 0.231
| Mean |  | 0.238 | 0.2335 | 0.23183 | 0.25 | 0.24533 | 0.22967 | 0.243 | 0.239 | 0.22693 | 0.20997 | 0.22237 | 0.23267 | 0.23448 | 0.26333 | 0.255 | 0.28803 | 0.22 | 0.24133 | 0.2525 | 0.24167 | 0.22867 | 0.23667 | 0.2065 | 0.25667 | 0.23483
| Median |  | 0.2375 | 0.2325 | 0.2315 | 0.25 | 0.2455 | 0.229 | 0.2435 | 0.2385 | 0.2275 | 0.20969 | 0.22118 | 0.232 | 0.23394 | 0.26 | 0.255 | 0.28656 | 0.22 | 0.241 | 0.253 | 0.24 | 0.231 | 0.2361 | 0.207 | 0.26 | 0.231
| Std Dev. |  | 0.0026077 | 0.0057879 | 0.0030605 | 0 | 0.0022509 | 0.0019664 | 0.0035777 | 0.0023664 | 0.0031136 | 0.0013049 | 0.0027429 | 0.0024221 | 0.0038382 | 0.005164 | 0.0054772 | 0.0037618 | 0 | 0.0026583 | 0.007287 | 0.0040825 | 0.0090701 | 0.003696 | 0.0036742 | 0.005164 | 0.01042
| Rel.Std.Dev. |  | 1.096% | 2.479% | 1.32% | 0% | 0.9175% | 0.8562% | 1.472% | 0.9901% | 1.372% | 0.6215% | 1.233% | 1.041% | 1.637% | 1.961% | 2.148% | 1.306% | 0% | 1.102% | 2.886% | 1.689% | 3.967% | 1.562% | 1.779% | 2.012% | 4.437%
| PDM3 |  | 0.1979% | -1.697% | -2.398% | 5.25% | 3.285% | -3.31% | 2.303% | 0.6189% | -4.461% | -11.6% | -6.382% | -2.047% | -1.283% | 10.86% | 7.355% | 21.26% | -7.38% | 1.601% | 6.302% | 1.742% | -3.731% | -0.3638% | -13.06% | 8.057% | -1.135%
| Z-Score (Absolute) |  | 0.11 | 0.27 | 0.41 | 1.12 | 0.73 | 0.59 | 0.53 | 0.2 | 0.82 | 2.25 | 1.2 | 0.34 | 0.18 | 2.25 | 1.55 | 4.33 | 1.4 | 0.39 | 1.33 | 0.42 | 0.67 | 0 | 2.54 | 1.69 | 0.15
| Analytical results for Tl in OREAS 505b (Certified Value 0.54 ppm) | [28 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.26 | 1.27 | 1.29 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g |  |  |  |  | 3
| 1 | 1 | 0.65 | 0.5 | 0.52 | 0.58 | 0.51 | 0.55 | 0.51 | 0.58 | 0.55 | 0.57768 | 0.64 | 1.4 | < 5 | 0.53 | 0.49633 | 0.24 | 0.49 | 0.53 | < 10 | 0.54 | 0.57 | 0.55 | 0.52 |  |  |  |  | 1
| 1 | 2 | 0.65 | 0.5 | 0.52 | 0.58 | 0.49 | 0.57 | 0.5 | 0.57 | 0.56 | 0.583 | 0.52 | 1.04 | < 5 | 0.53 | 0.52094 | 0.25 | 0.46 | 0.54 | < 10 | 0.54 | 0.551 | 0.55 | 0.54 |  |  |  |  | 33
| 1 | 3 | 0.64 | 0.5 | 0.52 | 0.58 | 0.49 | 0.57 | 0.53 | 0.56 | 0.55 | 0.5624 | 0.56 | 2.18 | < 5 | 0.53 | 0.53246 | 0.27 | 0.47 | 0.54 | < 10 | 0.58 | 0.571 | 0.54 | 0.53 |  |  |  |  | 16
| 1 | 4 | 0.67 | 0.6 | 0.52 | 0.57 | 0.49 | 0.56 | 0.61 | 0.57 | 0.56 | 0.5699 | 0.54 | 1.18 | < 5 | 0.5 | 0.51471 | 0.25 | 0.49 | 0.52 | < 10 | 0.57 | 0.576 | 0.54 | 0.53 |  |  |  |  | 0.54266
| 1 | 5 | 0.63 | 0.6 | 0.51 | 0.57 | 0.48 | 0.55 | 0.61 | 0.61 | 0.55 | 0.56593 | 0.55 | 1.17 | < 5 | 0.56 | 0.50443 | 0.25 | 0.49 | 0.54 | < 10 | 0.58 | 0.574 | 0.54 | 0.55 |  |  |  |  | 115
| 1 | 6 | 0.65 | 0.6 | 0.51 | 0.56 | 0.49 | 0.54 | 0.59 | 0.53 | 0.55 | 0.56856 | 0.49 | 0.57 | < 5 | 0.52 | 0.52533 | 0.25 | 0.49 | 0.56 | < 10 | 0.58 | 0.562 | 0.54 | 0.55
| Mean |  | 0.64833 | 0.55 | 0.51667 | 0.57333 | 0.49167 | 0.55667 | 0.55833 | 0.57 | 0.55333 | 0.57125 | 0.55 | 1.2567 |  | 0.52833 | 0.5157 | 0.25167 | 0.48167 | 0.53833 |  | 0.565 | 0.56733 | 0.54333 | 0.53667
| Median |  | 0.65 | 0.55 | 0.52 | 0.575 | 0.49 | 0.555 | 0.56 | 0.57 | 0.55 | 0.56923 | 0.545 | 1.175 |  | 0.53 | 0.51782 | 0.25 | 0.49 | 0.54 |  | 0.575 | 0.5705 | 0.54 | 0.535
| Std Dev. |  | 0.013292 | 0.054772 | 0.005164 | 0.008165 | 0.0098319 | 0.012111 | 0.050761 | 0.026077 | 0.005164 | 0.0076807 | 0.050596 | 0.53008 |  | 0.019408 | 0.013448 | 0.0098319 | 0.013292 | 0.013292 |  | 0.019748 | 0.009331 | 0.005164 | 0.012111
| Rel.Std.Dev. |  | 2.05% | 9.959% | 0.9995% | 1.424% | 2% | 2.176% | 9.092% | 4.575% | 0.9332% | 1.345% | 9.199% | 42.18% |  | 3.673% | 2.608% | 3.907% | 2.76% | 2.469% |  | 3.495% | 1.645% | 0.9504% | 2.257%
| PDM3 |  | 19.47% | 1.353% | -4.79% | 5.652% | -9.397% | 2.581% | 2.888% | 5.038% | 1.967% | 5.268% | 1.353% | 131.6% |  | -2.64% | -4.968% | -53.62% | -11.24% | -0.7972% |  | 4.117% | 4.547% | 0.1242% | -1.104%
| Z-Score (Absolute) |  | 3.43 | NA | 1.41 | 0.67 | 2.33 | 0.06 | 0.12 | 0.55 | 0.06 | 0.6 | 0.18 | 25.81 | 71.54 | 0.98 | 1.45 | 11.16 | 2.7 | 0.61 | 163.49 | 0.37 | 0.45 | 0.43 | 0.67
Results from laboratory 1.02 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Tm in OREAS 505b (Indicative Value 0.12 ppm) | [28 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  |  |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.11 |  |  |  |  | ppm
|  |  | 1.01: AR*MS; 1.02: AR*MS; 1.03: AR*MS |  |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.2g |  |  |  |  | 2
| 1 | 1 | 1.01: 0.118; 1.02: 0.1013; 1.03: 0.12135 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.117; 1.02: 0.1076; 1.03: 0.126 |  |  |  |  | 8
| 1 | 3 | 1.01: 0.12; 1.02: 0.1021; 1.03: 0.1213 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.116; 1.02: 0.1112; 1.03: 0.11985 |  |  |  |  | 0.11622
| 1 | 5 | 1.01: 0.12; 1.02: 0.1199; 1.03: 0.12173 |  |  |  |  | 14
| 1 | 6 | 1.01: 0.117; 1.02: 0.1111; 1.03: 0.12045
| Mean |  | 1.01: 0.118; 1.02: 0.10887; 1.03: 0.12178
| Median |  | 1.01: 0.1175; 1.02: 0.10935; 1.03: 0.12133
| Std Dev. |  | 1.01: 0.0016733; 1.02: 0.0068821; 1.03: 0.0021769
| Rel.Std.Dev. |  | 1.01: 1.418%; 1.02: 6.322%; 1.03: 1.788%
| PDM3 |  | 1.01: 1.535%; 1.02: -6.324%; 1.03: 4.788%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 1.63; 1.03: 0.67
| Analytical results for U in OREAS 505b (Certified Value 2.96 ppm) | [28 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.26 | 1.27 | 1.29 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | [6 empty] | 3
| 1 | 1 | 3.03 | 2.7 | 2.7 | 3.22 | 2.45 | 3.04 | 3.48 | 3.56 | 3.176 | 2.6162 | 2.9747 | 3.09 | 3.03 | 3.1065 | 2.98 | 2.49 | 3.29 | 2.9 | 3.196 | 3 | 2.88 | [6 empty] | 1
| 1 | 2 | 3.02 | 3 | 2.6 | 3.42 | 2.41 | 3.63 | 3.14 | 3.4 | 3.509 | 2.6722 | 2.8146 | 2.76 | 2.73 | 3.0851 | 3.22 | 2.53 | 2.87 | 3.2 | 3.21 | 3.2 | 3.17 | [6 empty] | 35
| 1 | 3 | 3.02 | 2.7 | 2.6 | 2.79 | 2.81 | 2.8 | 3.76 | 2.94 | 3.036 | 2.6874 | 2.5856 | 3.02 | 2.77 | 3.2641 | 3.1 | 2.48 | 2.87 | 3.2 | 3.155 | 3.1 | 3.23 | [6 empty] | 16
| 1 | 4 | 3.14 | 3 | 2.6 | 2.46 | 2.83 | 2.98 | 3.29 | 3.17 | 2.681 | 3.0805 | 2.7058 | 2.93 | 2.87 | 3.1848 | 2.85 | 3.13 | 2.88 | 2.7 | 3.216 | 3.1 | 2.94 | [6 empty] | 2.9645
| 1 | 5 | 3.14 | 3 | 2.6 | 2.93 | 2.37 | 3.46 | 2.62 | 3.15 | 2.711 | 3.2098 | 3.0573 | 2.63 | 3.07 | 3.107 | 3.08 | 2.58 | 3.22 | 3 | 3.218 | 3 | 3.18 | [6 empty] | 116
| 1 | 6 | 3.23 | 2.8 | 2.6 | 2.91 | 2.64 | 2.77 | 3.02 | 3.32 | 3.087 | 2.6261 | 3.104 | 2.86 | 2.65 | 3.2209 | 3.06 | 2.77 | 2.68 | 3 | 3.171 | 3 | 3.08
| Mean |  | 3.0967 | 2.8667 | 2.6167 | 2.955 | 2.585 | 3.1133 | 3.2183 | 3.2567 | 3.0333 | 2.8154 | 2.8737 | 2.8817 | 2.8533 | 3.1614 | 3.0483 | 2.6633 | 2.9683 | 3 | 3.1943 | 3.0667 | 3.08
| Median |  | 3.085 | 2.9 | 2.6 | 2.92 | 2.545 | 3.01 | 3.215 | 3.245 | 3.0615 | 2.6798 | 2.8946 | 2.895 | 2.82 | 3.1459 | 3.07 | 2.555 | 2.875 | 3 | 3.203 | 3.05 | 3.125
| Std Dev. |  | 0.086872 | 0.15055 | 0.040825 | 0.33507 | 0.20433 | 0.35393 | 0.39255 | 0.21713 | 0.30904 | 0.26009 | 0.20574 | 0.1694 | 0.16848 | 0.072695 | 0.12432 | 0.25201 | 0.23541 | 0.18974 | 0.025959 | 0.08165 | 0.14156
| Rel.Std.Dev. |  | 2.805% | 5.252% | 1.56% | 11.34% | 7.904% | 11.37% | 12.2% | 6.667% | 10.19% | 9.238% | 7.16% | 5.879% | 5.905% | 2.299% | 4.078% | 9.462% | 7.931% | 6.325% | 0.8127% | 2.662% | 4.596%
| PDM3 |  | 4.458% | -3.301% | -11.73% | -0.3209% | -12.8% | 5.02% | 8.562% | 9.855% | 2.321% | -5.031% | -3.064% | -2.795% | -3.75% | 6.641% | 2.827% | -10.16% | 0.1289% | 1.197% | 7.752% | 3.446% | 3.896%
| Z-Score (Absolute) |  | 0.52 | 0.71 | 2.05 | 0.24 | 2.22 | 0.6 | 1.17 | 1.37 | 0.18 | 0.99 | 0.67 | 0.63 | 0.78 | 0.86 | 0.26 | 1.8 | 0.17 | 0 | 1.04 | 0.36 | 0.43
| Analytical results for V in OREAS 505b (Certified Value 71 ppm) | [28 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.27 | 1.29 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES |  | 0
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g |  | 1
| 1 | 1 | 73 | 66 | 70 | 75 | 72 | 73 | 73 | 69 | 72.3 | 68.0679 | 71.5266 | 72.9 | 74.67 | 71.2124 | 69 | 74 | 67.4234 | 72 | 68 | 75 | 71.4 | 68 | 69.85 | 71 | 78 | 70 |  | 1
| 1 | 2 | 72 | 67 | 70 | 75 | 72 | 72 | 73 | 70 | 73.4 | 69.0041 | 69.911 | 70.4 | 72.11 | 72.4552 | 69.8 | 75 | 70.8237 | 68 | 68 | 77 | 72.2 | 68 | 69.81 | 70 | 79 | 69 |  | 17
| 1 | 3 | 72 | 68 | 70 | 75 | 71 | 72 | 73 | 70 | 74.8 | 68.9535 | 69.0953 | 73.5 | 72.92 | 72.1111 | 68.6 | 76 | 68.7109 | 71 | 68 | 75 | 72 | 72 | 70.34 | 70 | 80 | 68 |  | 16
| 1 | 4 | 73 | 68 | 70 | 76 | 72 | 72 | 72 | 70 | 73.3 | 68.4691 | 70.6578 | 73.7 | 73.67 | 71.7096 | 69.1 | 73 | 69.2715 | 70 | 68 | 71 | 72 | 71 | 69.87 | 70 | 80 | 68 |  | 71.1402
| 1 | 5 | 71 | 69 | 70 | 76 | 71 | 72 | 71 | 70 | 75.1 | 68.8703 | 69.1103 | 71.3 | 72 | 75.2946 | 68.6 | 75 | 71.5415 | 71 | 69 | 71 | 72.2 | 72 | 71.1 | 69 | 78 | 68 |  | 117
| 1 | 6 | 73 | 68 | 70 | 76 | 71 | 72 | 71 | 71 | 74.4 | 70.7887 | 69.2168 | 73.1 | 71.48 | 72.5891 | 69.6 | 74 | 69.6016 | 72 | 66 | 74 | 73.3 | 72 | 71.07 | 69 | 81 | 69
| Mean |  | 72.3333 | 67.6667 | 70 | 75.5 | 71.5 | 72.1667 | 72.1667 | 70 | 73.8833 | 69.0256 | 69.9196 | 72.4833 | 72.8083 | 72.562 | 69.1167 | 74.5 | 69.5621 | 70.6667 | 67.8333 | 73.8333 | 72.1833 | 70.5 | 70.34 | 69.8333 | 79.3333 | 68.6667
| Median |  | 72.5 | 68 | 70 | 75.5 | 71.5 | 72 | 72.5 | 70 | 73.9 | 68.9119 | 69.5639 | 73 | 72.515 | 72.2832 | 69.05 | 74.5 | 69.4366 | 71 | 68 | 74.5 | 72.1 | 71.5 | 70.105 | 70 | 79.5 | 68.5
| Std Dev. |  | 0.8165 | 1.0328 | 0 | 0.54772 | 0.54772 | 0.40825 | 0.98319 | 0.63246 | 1.0647 | 0.93435 | 0.99558 | 1.3273 | 1.1938 | 1.4304 | 0.49967 | 1.0488 | 1.4761 | 1.5055 | 0.98319 | 2.4014 | 0.62102 | 1.9748 | 0.60867 | 0.75277 | 1.2111 | 0.8165
| Rel.Std.Dev. |  | 1.129% | 1.526% | 0% | 0.7255% | 0.766% | 0.5657% | 1.362% | 0.9035% | 1.441% | 1.354% | 1.424% | 1.831% | 1.64% | 1.971% | 0.7229% | 1.408% | 2.122% | 2.13% | 1.449% | 3.252% | 0.8603% | 2.801% | 0.8653% | 1.078% | 1.527% | 1.189%
| PDM3 |  | 1.677% | -4.883% | -1.603% | 6.128% | 0.5058% | 1.443% | 1.443% | -1.603% | 3.856% | -2.972% | -1.716% | 1.888% | 2.345% | 1.999% | -2.844% | 4.723% | -2.218% | -0.6656% | -4.648% | 3.786% | 1.466% | -0.8999% | -1.125% | -1.837% | 11.52% | -3.477%
| Z-Score (Absolute) |  | 0.59 | 1.6 | 0.51 | 2.07 | 0.2 | 0.51 | 0.51 | 0.51 | 1.31 | 0.96 | 0.55 | 0.66 | 0.81 | 0.69 | 0.92 | 1.6 | 0.71 | 0.2 | 1.52 | 1.29 | 0.52 | 0.27 | 0.35 | 0.59 | 3.87 | 1.13
| Analytical results for W in OREAS 505b (Certified Value 5.81 ppm) | [28 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.27 | 1.29 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g |  |  |  |  | 2
| 1 | 1 | 7.2 | 4.3 | 4.3 | 5.9 | 5.73 | 5.72 | 5.43 | 6.16 | 5.13 | 6.2305 | < 10 | 6.1 | < 5 | 5.4 | 6.3218 | 4.96 | 5.88 | 5.64 | < 10 | 3.5 | 7.5 | 6.3 | < 10 |  |  |  |  | 1
| 1 | 2 | 7.1 | 4.3 | 4.1 | 5.7 | 5.6 | 5.75 | 5.93 | 6.81 | 5.06 | 5.6122 | < 10 | 6.04 | < 5 | 5.5 | 6.0114 | 4.79 | 5.3 | 5.92 | < 10 | 3.63 | 7.2 | 6.9 | < 10 |  |  |  |  | 37
| 1 | 3 | 7.2 | 4.4 | 4.3 | 6.3 | 5.81 | 5.79 | 5.5 | 6.26 | 4.91 | 5.6546 | < 10 | 6.34 | < 5 | 5.1 | 6.1515 | 5.04 | 5.32 | 5.56 | < 10 | 3.66 | 7.2 | 6.2 | < 10 |  |  |  |  | 16
| 1 | 4 | 7.4 | 4.3 | 4.3 | 5.7 | 5.71 | 5.72 | 6.09 | 6.24 | 5.3 | 5.2623 | < 10 | 5.99 | < 5 | 5.1 | 6.0685 | 4.86 | 5.57 | 5.36 | < 10 | 3.56 | 7.5 | 6 | < 10 |  |  |  |  | 5.8129
| 1 | 5 | 7.3 | 4.3 | 4.1 | 5.8 | 5.9 | 5.83 | 5.99 | 5.83 | 5.03 | 4.6438 | < 10 | 6.08 | < 5 | 5.5 | 6.2771 | 4.89 | 5.44 | 6.13 | < 10 | 3.54 | 7 | 6.3 | < 10 |  |  |  |  | 118
| 1 | 6 | 7.2 | 4.2 | 4.1 | 6 | 5.71 | 6.25 | 5.99 | 5.72 | 5 | 5.268 | < 10 | 5.79 | < 5 | 5.4 | 6.0992 | 4.92 | 5.49 | 5.89 | < 10 | 3.53 | 7.2 | 6.7 | < 10
| Mean |  | 7.2333 | 4.3 | 4.2 | 5.9 | 5.7433 | 5.8433 | 5.8217 | 6.17 | 5.0717 | 5.4452 |  | 6.0567 |  | 5.3333 | 6.1549 | 4.91 | 5.5 | 5.75 |  | 3.57 | 7.2667 | 6.4
| Median |  | 7.2 | 4.3 | 4.2 | 5.85 | 5.72 | 5.77 | 5.96 | 6.2 | 5.045 | 5.4401 |  | 6.06 |  | 5.4 | 6.1254 | 4.905 | 5.465 | 5.765 |  | 3.55 | 7.2 | 6.3
| Std Dev. |  | 0.10328 | 0.063246 | 0.10954 | 0.22804 | 0.10191 | 0.20373 | 0.28188 | 0.38492 | 0.13318 | 0.5283 |  | 0.17806 |  | 0.18619 | 0.1216 | 0.08579 | 0.21232 | 0.28043 |  | 0.061968 | 0.19664 | 0.33466
| Rel.Std.Dev. |  | 1.428% | 1.471% | 2.608% | 3.865% | 1.774% | 3.487% | 4.842% | 6.239% | 2.626% | 9.702% |  | 2.94% |  | 3.491% | 1.976% | 1.747% | 3.86% | 4.877% |  | 1.736% | 2.706% | 5.229%
| PDM3 |  | 24.44% | -26.03% | -27.75% | 1.499% | -1.196% | 0.524% | 0.1513% | 6.144% | -12.75% | -6.325% |  | 4.194% |  | -8.25% | 5.884% | -15.53% | -5.382% | -1.082% |  | -38.58% | 25.01% | 10.1%
| Z-Score (Absolute) |  | 2.34 | 1.62 | 1.75 | 0.54 | 0.33 | 0.46 | 0.43 | 0.9 | 0.58 | 0.07 | 0.67 | 0.75 | 4.05 | 0.22 | 0.88 | 0.8 | 0 | 0.34 | 0.67 | 2.6 | 2.38 | 1.21 | 0.67
| Analytical results for Y in OREAS 505b (Certified Value 10.6 ppm) | [28 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | [7 empty] | 1
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | [7 empty] | 2
| 1 | 1 | 11.2 | 10.02 | 10.75 | 10.45 | 11.65 | 10.6 | 10.64 | 10.7358 | 10.4377 | 11.4 | 10.6 | 9.9706 | 10.76 | 9.04 | 10.7 | 10.6 | 10.576 | 10.1 | 10.4 | 10.4 | [7 empty] | 1
| 1 | 2 | 11.4 | 9.77 | 10.85 | 10.4 | 11 | 10.25 | 10.78 | 10.7744 | 10.3147 | 9.5 | 10.6 | 10.5493 | 10.61 | 9.19 | 11.55 | 10.8 | 10.636 | 10 | 10.9 | 10.4 | [7 empty] | 38
| 1 | 3 | 11.5 | 9.97 | 11.35 | 10.65 | 11.35 | 10.4 | 10.77 | 10.8703 | 10.283 | 9.64 | 11 | 10.1265 | 10.41 | 9.07 | 10.75 | 10.8 | 10.638 | 9.9 | 10.6 | 9.6 | [7 empty] | 16
| 1 | 4 | 11.1 | 10.05 | 10.8 | 10.45 | 10.8 | 10.3 | 10.57 | 10.6344 | 10.4243 | 10.3 | 10.7 | 10.289 | 10.32 | 8.85 | 10.25 | 10.8 | 10.604 | 9.9 | 11 | 9.7 | [7 empty] | 10.5684
| 1 | 5 | 11.2 | 10.12 | 11.2 | 10.25 | 10.55 | 10.6 | 10.96 | 10.7938 | 10.2735 | 9.41 | 10.9 | 10.662 | 10.32 | 9.62 | 10.2 | 10.8 | 10.728 | 9.8 | 10.7 | 9.8 | [7 empty] | 119
| 1 | 6 | 11 | 9.98 | 10.85 | 10 | 10.35 | 9.62 | 10.86 | 11.1329 | 10.3123 | 9.83 | 10.8 | 10.4855 | 10.56 | 9.24 | 10.7 | 11 | 10.809 | 9.8 | 10.9 | 10.1
| Mean |  | 11.2333 | 9.985 | 10.9667 | 10.3667 | 10.95 | 10.295 | 10.7633 | 10.8236 | 10.3409 | 10.0133 | 10.7667 | 10.3472 | 10.4967 | 9.1683 | 10.6917 | 10.8 | 10.6652 | 9.9167 | 10.75 | 10
| Median |  | 11.2 | 10 | 10.85 | 10.425 | 10.9 | 10.35 | 10.775 | 10.7841 | 10.3135 | 9.735 | 10.75 | 10.3873 | 10.485 | 9.13 | 10.7 | 10.8 | 10.637 | 9.9 | 10.8 | 9.95
| Std Dev. |  | 0.18619 | 0.11845 | 0.24631 | 0.22061 | 0.48888 | 0.36187 | 0.14208 | 0.17008 | 0.071757 | 0.74907 | 0.1633 | 0.26572 | 0.17671 | 0.25965 | 0.4852 | 0.12649 | 0.087083 | 0.1169 | 0.22583 | 0.35214
| Rel.Std.Dev. |  | 1.657% | 1.186% | 2.246% | 2.128% | 4.465% | 3.515% | 1.32% | 1.571% | 0.6939% | 7.481% | 1.517% | 2.568% | 1.683% | 2.832% | 4.538% | 1.171% | 0.8165% | 1.179% | 2.101% | 3.521%
| PDM3 |  | 6.291% | -5.52% | 3.768% | -1.909% | 3.611% | -2.587% | 1.844% | 2.414% | -2.153% | -5.252% | 1.876% | -2.094% | -0.679% | -13.25% | 1.166% | 2.191% | 0.9154% | -6.167% | 1.718% | -5.379%
| Z-Score (Absolute) |  | 1.82 | 1.67 | 1.08 | 0.6 | 1.03 | 0.8 | 0.51 | 0.68 | 0.67 | 1.59 | 0.52 | 0.65 | 0.24 | 3.95 | 0.31 | 0.61 | 0.24 | 1.86 | 0.47 | 1.62
| Analytical results for Yb in OREAS 505b (Certified Value 0.73 ppm) | [28 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [7 empty] | 1
| No | No | 1.01: 1.09; 1.04: 1.10; 1.05: 1.11; 1.06: 1.17; 1.07: 1.29 | [7 empty] | ppm
|  |  | 1.01: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*OES/MS; 1.07: AR*OES/MS | [7 empty] | 2
|  |  | 1.01: 0.5g; 1.04: 0.5g; 1.05: 0.2g; 1.06: 0.25g; 1.07: 0.25g | [7 empty] | 3
| 1 | 1 | 1.01: 0.726; 1.04: 0.6631; 1.05: 0.79586; 1.06: 0.7; 1.07: 0.8 | [7 empty] | 1
| 1 | 2 | 1.01: 0.728; 1.04: 0.6921; 1.05: 0.7792; 1.06: 0.7; 1.07: 0.8 | [7 empty] | 39
| 1 | 3 | 1.01: 0.721; 1.04: 0.6717; 1.05: 0.7613; 1.06: 0.7; 1.07: 0.7 | [7 empty] | 16
| 1 | 4 | 1.01: 0.717; 1.04: 0.6969; 1.05: 0.75884; 1.06: 0.7; 1.07: 0.7 | [7 empty] | 0.73257
| 1 | 5 | 1.01: 0.741; 1.04: 0.7656; 1.05: 0.79054; 1.06: 0.7; 1.07: 0.8 | [7 empty] | 120
| 1 | 6 | 1.01: 0.767; 1.04: 0.7004; 1.05: 0.78233; 1.06: 0.7; 1.07: 0.8
| Mean |  | 1.01: 0.73333; 1.04: 0.6983; 1.05: 0.77801; 1.06: 0.7; 1.07: 0.76667
| Median |  | 1.01: 0.727; 1.04: 0.6945; 1.05: 0.78077; 1.06: 0.7; 1.07: 0.8
| Std Dev. |  | 1.01: 0.018403; 1.04: 0.036113; 1.05: 0.015115; 1.06: 1.2162e-16; 1.07: 0.05164
| Rel.Std.Dev. |  | 1.01: 2.509%; 1.04: 5.172%; 1.05: 1.943%; 1.06: 1.737e-14%; 1.07: 6.736%
| PDM3 |  | 1.01: 0.1042%; 1.04: -4.678%; 1.05: 6.203%; 1.06: -4.446%; 1.07: 4.654%
| Z-Score (Absolute) |  | 1.01: 0; 1.04: 0.71; 1.05: 0.9; 1.06: 0.67; 1.07: 0.67
| Analytical results for Zn in OREAS 505b (Certified Value 95 ppm) | [28 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.26 | 1.27 | 1.29 | 1.30 | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | 0
|  |  | 01g | 0.5g | 15g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 0.5g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.25g | 1
| 1 | 1 | 93 | 88 | 95.7 | 98 | 97 | 97 | 96 | 91 | 94.1 | 94.6323 | 97.2176 | 95.1 | 98.93 | 89.25 | 96.8 | 90 | 91.6932 | 95 | 95 | 97 | 94.2 | 92 | 93.313 | 89.9491 | 90 | 102 | 94 | 1
| 1 | 2 | 94 | 91 | 92.2 | 99 | 99 | 95 | 95 | 91 | 93.4 | 93.2629 | 95.3428 | 93.4 | 99.21 | 92.67 | 94.6 | 91 | 93.3261 | 93 | 95 | 100 | 95.6 | 93.9 | 92.047 | 89.9497 | 89 | 106 | 97 | 18
| 1 | 3 | 94 | 91 | 95.7 | 98 | 97 | 96 | 96 | 92 | 93.7 | 96.36 | 95.1635 | 101 | 96.28 | 91.1 | 93.4 | 92 | 90.2817 | 93 | 95 | 97 | 95.4 | 97.2 | 96.17 | 90.024 | 89 | 98 | 98 | 16
| 1 | 4 | 92 | 92 | 91.5 | 98 | 97 | 95 | 94 | 92 | 91 | 96.5518 | 95.9846 | 98.9 | 98.18 | 93.71 | 93 | 91 | 95.1689 | 93 | 94 | 94 | 95.7 | 96.2 | 92.464 | 89.9493 | 89 | 103 | 94 | 94.5246
| 1 | 5 | 93 | 93 | 94.8 | 98 | 97 | 94 | 93 | 90 | 95.2 | 95.8414 | 95.1246 | 94.7 | 96.238 | 92.08 | 95.9 | 93 | 95.9568 | 92 | 96 | 93 | 95.8 | 97.3 | 94.395 | 90.0454 | 88 | 104 | 95 | 121
| 1 | 6 | 96 | 92 | 92.7 | 98 | 95 | 95 | 92 | 91 | 94.9 | 92.5642 | 101.1351 | 97.7 | 93.57 | 86.95 | 96.9 | 90 | 97.255 | 91 | 93 | 95 | 97.1 | 98.8 | 94.985 | 89.9255 | 88 | 108 | 96
| Mean |  | 93.6667 | 91.1667 | 93.7667 | 98.1667 | 97 | 95.3333 | 94.3333 | 91.1667 | 93.7167 | 94.8688 | 96.6614 | 96.8 | 97.068 | 90.96 | 95.1 | 91.1667 | 93.9469 | 92.8333 | 94.6667 | 96 | 95.6333 | 95.9 | 93.8957 | 89.9738 | 88.8333 | 103.5 | 95.6667
| Median |  | 93.5 | 91.5 | 93.75 | 98 | 97 | 95 | 94.5 | 91 | 93.9 | 95.2368 | 95.6637 | 96.4 | 97.23 | 91.59 | 95.25 | 91 | 94.2475 | 93 | 95 | 96 | 95.65 | 96.7 | 93.854 | 89.9495 | 89 | 103.5 | 95.5
| Std Dev. |  | 1.3663 | 1.7224 | 1.8587 | 0.40825 | 1.2649 | 1.0328 | 1.633 | 0.75277 | 1.4986 | 1.6702 | 2.3292 | 2.8886 | 2.1372 | 2.4805 | 1.6923 | 1.169 | 2.6599 | 1.3292 | 1.0328 | 2.5298 | 0.92664 | 2.5044 | 1.5758 | 0.048513 | 0.75277 | 3.4496 | 1.633
| Rel.Std.Dev. |  | 1.459% | 1.889% | 1.982% | 0.4159% | 1.304% | 1.083% | 1.731% | 0.8257% | 1.599% | 1.761% | 2.41% | 2.984% | 2.202% | 2.727% | 1.78% | 1.282% | 2.831% | 1.432% | 1.091% | 2.635% | 0.969% | 2.611% | 1.678% | 0.05392% | 0.8474% | 3.333% | 1.707%
| PDM3 |  | -0.9076% | -3.552% | -0.8018% | 3.853% | 2.619% | 0.8556% | -0.2023% | -3.552% | -0.8547% | 0.3642% | 2.261% | 2.407% | 2.691% | -3.771% | 0.6088% | -3.552% | -0.6111% | -1.789% | 0.1503% | 1.561% | 1.173% | 1.455% | -0.6653% | -4.814% | -6.021% | 9.495% | 1.208%
| Z-Score (Absolute) |  | 0.51 | 1.77 | 0.46 | 1.77 | 1.18 | 0.34 | 0.17 | 1.77 | 0.48 | 0.1 | 1.01 | 1.08 | 1.21 | 1.87 | 0.22 | 1.77 | 0.36 | 0.93 | 0 | 0.67 | 0.49 | 0.62 | 0.39 | 2.37 | 2.95 | 4.47 | 0.51
| Analytical results for Zr in OREAS 505b (Certified Value 8.56 ppm) | [28 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.24 | 1.27 | 1.29 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | [6 empty] | 2
|  |  | 01g | 0.25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.2g | 0.5g | 01g | 0.25g | 01g | 0.2g | 0.5g | 0.5g | 0.5g | 0.25g | 0.25g | 0.25g | 0.25g | [6 empty] | 3
| 1 | 1 | 9.2 | 9.1 | 9.6 | 8.3 | 10.1 | 8.4 | 8.26 | 8.416 | 8.8106 | 8.1 | 8.5 | 7.6 | 7.6706 | 7.8 | 7.7 | 8.1 | 10.7 | 9.6 | 11 | 8.1 | 9.4 | [6 empty] | 1
| 1 | 2 | 9.6 | 9.2 | 9.7 | 8.4 | 9.8 | 8.1 | 8.18 | 6.9746 | 8.9646 | 7 | 8.1 | 7.8 | 8.2943 | 6.8 | 7.9 | 8.5 | 10.6 | 9.38 | 10.7 | 8.7 | 9.6 | [6 empty] | 19
| 1 | 3 | 9.7 | 9.2 | 9.8 | 8.5 | 10.1 | 8.3 | 8.05 | 7.9786 | 8.7849 | 6.8 | 8.6 | 7.9 | 8.0125 | 8.1 | 7.8 | 8.2 | 10.5 | 9.58 | 10.5 | 8.2 | 9.7 | [6 empty] | 16
| 1 | 4 | 9.6 | 9.6 | 9.3 | 8.3 | 9 | 8.3 | 8.12 | 7.2333 | 9.0148 | 7.5 | 8.5 | 7.8 | 8.5397 | 7.8 | 7.8 | 7.8 | 11 | 9.62 | 10.6 | 8.3 | 9.8 | [6 empty] | 8.5573
| 1 | 5 | 9.7 | 9.8 | 9.8 | 8.3 | 8.7 | 8.4 | 8.09 | 7.2059 | 8.6848 | 7.3 | 8.9 | 7.9 | 8.8169 | 7.6 | 8.3 | 7.8 | 10.7 | 9.63 | 10.5 | 8.2 | 9.8 | [6 empty] | 122
| 1 | 6 | 9.6 | 9.7 | 9.7 | 7.9 | 8.6 | 7.8 | 8.19 | 8.4649 | 8.4378 | 7 | 8.6 | 7.9 | 7.8707 | 7.9 | 7.8 | 8.3 | 11 | 9.72 | 10.4 | 8.8 | 10.3
| Mean |  | 9.5667 | 9.4333 | 9.65 | 8.2833 | 9.3833 | 8.2167 | 8.1483 | 7.7122 | 8.7829 | 7.2833 | 8.5333 | 7.8167 | 8.2008 | 7.6667 | 7.8833 | 8.1167 | 10.75 | 9.5883 | 10.6167 | 8.3833 | 9.7667
| Median |  | 9.6 | 9.4 | 9.7 | 8.3 | 9.4 | 8.3 | 8.15 | 7.6059 | 8.7977 | 7.15 | 8.55 | 7.85 | 8.1534 | 7.8 | 7.8 | 8.15 | 10.7 | 9.61 | 10.55 | 8.25 | 9.75
| Std Dev. |  | 0.18619 | 0.30111 | 0.18708 | 0.20412 | 0.6969 | 0.23166 | 0.076267 | 0.65766 | 0.20798 | 0.47081 | 0.2582 | 0.1169 | 0.43099 | 0.45461 | 0.2137 | 0.27869 | 0.20736 | 0.11286 | 0.2137 | 0.29269 | 0.30111
| Rel.Std.Dev. |  | 1.946% | 3.192% | 1.939% | 2.464% | 7.427% | 2.819% | 0.936% | 8.528% | 2.368% | 6.464% | 3.026% | 1.496% | 5.256% | 5.93% | 2.711% | 3.434% | 1.929% | 1.177% | 2.013% | 3.491% | 3.083%
| PDM3 |  | 11.8% | 10.24% | 12.77% | -3.202% | 9.653% | -3.981% | -4.779% | -9.876% | 2.636% | -14.89% | -0.2804% | -8.655% | -4.167% | -10.41% | -7.876% | -5.15% | 25.62% | 12.05% | 24.07% | -2.033% | 14.13%
| Z-Score (Absolute) |  | 1.19 | 1.05 | 1.27 | 0.1 | 1 | 0.17 | 0.24 | 0.67 | 0.4 | 1.11 | 0.15 | 0.57 | 0.18 | 0.72 | 0.5 | 0.27 | 2.38 | 1.21 | 2.24 | 0 | 1.39
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C36: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C54: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C73: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C91: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C128: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C147: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C165: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C184: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C202: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C221: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C239: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C257: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C275: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C311: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C329: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C347: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C366: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C384: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C403: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C422: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C441: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C459: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C477: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C495: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C514: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C532: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C550: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C568: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C586: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C604: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C622: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C641: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C659: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C677: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C695: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C731: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C749: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C767: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C785: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C803: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C821: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C840: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C859: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1005: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1023: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1041: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1060: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1078: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1096: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1114: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1132: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1186: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 505b (Indicative Value 15.01 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.99 | 1
| 1 | 2 | 15.02 | 12
| Mean |  | 15.005 | 16
| Median |  | 15.005 | 15.005
| Std Dev. |  | 0.021213 | 18
| Rel.Std.Dev. |  | 0.1414%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 505b (Indicative Value 10 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 10 | 1
| 1 | 2 | 10 | 13
| Mean |  | 10 | 16
| Median |  | 10 | 10
| Std Dev. |  | 0 | 19
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 505b (Indicative Value 990 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 990 | 1
| 1 | 2 | 990 | 14
| Mean |  | 990 | 16
| Median |  | 990 | 990
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 505b (Indicative Value 2.62 wt.%) |  |  | Indicative
CaO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.61 | 1
| 1 | 2 | 2.62 | 15
| Mean |  | 2.615 | 16
| Median |  | 2.615 | 2.615
| Std Dev. |  | 0.0070711 | 21
| Rel.Std.Dev. |  | 0.2704%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cl in OREAS 505b (Indicative Value 40 ppm) |  |  | Indicative
Cl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 50 | 1
| 1 | 2 | 30 | 16
| Mean |  | 40 | 16
| Median |  | 40 | 40
| Std Dev. |  | 14.1421 | 22
| Rel.Std.Dev. |  | 35.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 505b (Indicative Value 20 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 17
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 505b (Indicative Value 55 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 60 | 1
| 1 | 2 | 50 | 18
| Mean |  | 55 | 16
| Median |  | 55 | 55
| Std Dev. |  | 7.0711 | 24
| Rel.Std.Dev. |  | 12.86%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 505b (Indicative Value 0.329 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.329 | 1
| 1 | 2 | 0.328 | 19
| Mean |  | 0.3285 | 16
| Median |  | 0.3285 | 0.3285
| Std Dev. |  | 0.00070711 | 25
| Rel.Std.Dev. |  | 0.2153%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 505b (Indicative Value 3.18 wt.%) |  |  | Indicative
Fe | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.18 | 1
| 1 | 2 | 3.18 | 12
| Mean |  | 3.18 | 16
| Median |  | 3.18 | 3.18
| Std Dev. |  | 0 | 18
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 505b (Indicative Value 3.96 wt.%) |  |  | Indicative
K2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.95 | 1
| 1 | 2 | 3.96 | 13
| Mean |  | 3.955 | 16
| Median |  | 3.955 | 3.955
| Std Dev. |  | 0.0070711 | 19
| Rel.Std.Dev. |  | 0.1788%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 505b (Indicative Value 1.37 wt.%) |  |  | Indicative
MgO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 1.37 | 1
| 1 | 2 | 1.37 | 14
| Mean |  | 1.37 | 16
| Median |  | 1.37 | 1.37
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 505b (Indicative Value 0.058 wt.%) |  |  | Indicative
MnO | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.058 | 1
| 1 | 2 | 0.058 | 15
| Mean |  | 0.058 | 16
| Median |  | 0.058 | 0.058
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 505b (Indicative Value 2.93 wt.%) |  |  | Indicative
Na2O | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.93 | 1
| 1 | 2 | 2.92 | 16
| Mean |  | 2.925 | 16
| Median |  | 2.925 | 2.925
| Std Dev. |  | 0.0070711 | 22
| Rel.Std.Dev. |  | 0.2417%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 505b (Indicative Value 50 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 17
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 23
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 505b (Indicative Value 0.087 wt.%) |  |  | Indicative
P | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.087 | 1
| 1 | 2 | 0.086 | 18
| Mean |  | 0.0865 | 16
| Median |  | 0.0865 | 0.0865
| Std Dev. |  | 0.00070711 | 24
| Rel.Std.Dev. |  | 0.8175%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 505b (Indicative Value 70 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 70 | 19
| Mean |  | 70 | 16
| Median |  | 70 | 70
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 505b (Indicative Value 0.555 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.554 | 1
| 1 | 2 | 0.555 | 23
| Mean |  | 0.5545 | 16
| Median |  | 0.5545 | 0.5545
| Std Dev. |  | 0.00070711 | 18
| Rel.Std.Dev. |  | 0.1275%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 505b (Indicative Value 66.51 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 66.52 | 1
| 1 | 2 | 66.49 | 13
| Mean |  | 66.505 | 16
| Median |  | 66.505 | 66.505
| Std Dev. |  | 0.021213 | 19
| Rel.Std.Dev. |  | 0.0319%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 505b (Indicative Value 20 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 14
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 505b (Indicative Value 340 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 340 | 1
| 1 | 2 | 340 | 15
| Mean |  | 340 | 16
| Median |  | 340 | 340
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 505b (Indicative Value 0.587 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.586 | 1
| 1 | 2 | 0.587 | 16
| Mean |  | 0.5865 | 16
| Median |  | 0.5865 | 0.5865
| Std Dev. |  | 0.00070711 | 22
| Rel.Std.Dev. |  | 0.1206%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 505b (Indicative Value 75 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 70 | 1
| 1 | 2 | 80 | 17
| Mean |  | 75 | 16
| Median |  | 75 | 75
| Std Dev. |  | 7.0711 | 23
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 505b (Indicative Value 115 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 110 | 1
| 1 | 2 | 120 | 18
| Mean |  | 115 | 16
| Median |  | 115 | 115
| Std Dev. |  | 7.0711 | 24
| Rel.Std.Dev. |  | 6.149%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 505b (Indicative Value 235 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 230 | 1
| 1 | 2 | 240 | 19
| Mean |  | 235 | 16
| Median |  | 235 | 235
| Std Dev. |  | 7.0711 | 25
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 505b (Indicative Value 1.47 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 1.48 | 1
| 1 | 2 | 1.45 | 21
| Mean |  | 1.465 | 16
| Median |  | 1.465 | 1.465
| Std Dev. |  | 0.021213 | 27
| Rel.Std.Dev. |  | 1.448%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 505b (Indicative Value 0.135 wt.%) |  |  | Indicative
C | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.14 | 1
| 1 | 2 | 0.13 | 23
| Mean |  | 0.135 | 16
| Median |  | 0.135 | 0.135
| Std Dev. |  | 0.0070711 | 29
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 505b (Indicative Value 0.47 wt.%) |  |  | Indicative
S | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.46 | 1
| 1 | 2 | 0.48 | 23
| Mean |  | 0.47 | 16
| Median |  | 0.47 | 0.47
| Std Dev. |  | 0.014142 | 29
| Rel.Std.Dev. |  | 3.009%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 505b (Indicative Value 1.3 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.3 | 25
| Mean |  | 1.3 | 16
| Median |  | 1.3 | 1.3
| Std Dev. |  | 0 | 31
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 505b (Indicative Value 9 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.8 | 1
| 1 | 2 | 9.2 | 13
| Mean |  | 9 | 16
| Median |  | 9 | 9
| Std Dev. |  | 0.28284 | 32
| Rel.Std.Dev. |  | 3.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 505b (Indicative Value 927 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 921 | 1
| 1 | 2 | 933 | 14
| Mean |  | 927 | 16
| Median |  | 927 | 927
| Std Dev. |  | 8.4853 | 33
| Rel.Std.Dev. |  | 0.9153%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 505b (Indicative Value 2.9 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 3 | 28
| Mean |  | 2.9 | 16
| Median |  | 2.9 | 2.9
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 4.877%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 505b (Indicative Value 0.84 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.76 | 1
| 1 | 2 | 0.92 | 29
| Mean |  | 0.84 | 16
| Median |  | 0.84 | 0.84
| Std Dev. |  | 0.11314 | 35
| Rel.Std.Dev. |  | 13.47%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 505b (Indicative Value 0.35 ppm) |  |  | Indicative
Cd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.3 | 1
| 1 | 2 | 0.4 | 30
| Mean |  | 0.35 | 16
| Median |  | 0.35 | 0.35
| Std Dev. |  | 0.070711 | 36
| Rel.Std.Dev. |  | 20.2%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 505b (Indicative Value 68 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 68.6 | 1
| 1 | 2 | 68.1 | 31
| Mean |  | 68.35 | 16
| Median |  | 68.35 | 68.35
| Std Dev. |  | 0.35355 | 37
| Rel.Std.Dev. |  | 0.5173%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 505b (Indicative Value 9.25 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.7 | 1
| 1 | 2 | 9.8 | 17
| Mean |  | 9.25 | 16
| Median |  | 9.25 | 9.25
| Std Dev. |  | 0.77782 | 38
| Rel.Std.Dev. |  | 8.409%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 505b (Indicative Value 70 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 69 | 1
| 1 | 2 | 70 | 18
| Mean |  | 69.5 | 16
| Median |  | 69.5 | 69.5
| Std Dev. |  | 0.70711 | 39
| Rel.Std.Dev. |  | 1.017%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 505b (Indicative Value 9.27 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 9.33 | 1
| 1 | 2 | 9.2 | 34
| Mean |  | 9.265 | 16
| Median |  | 9.265 | 9.265
| Std Dev. |  | 0.091924 | 40
| Rel.Std.Dev. |  | 0.9922%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 505b (Indicative Value 0.319 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.31 | 1
| 1 | 2 | 0.327 | 19
| Mean |  | 0.3185 | 16
| Median |  | 0.3185 | 0.3185
| Std Dev. |  | 0.012021 | 41
| Rel.Std.Dev. |  | 3.774%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 505b (Indicative Value 5.11 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.06 | 1
| 1 | 2 | 5.15 | 6
| Mean |  | 5.105 | 16
| Median |  | 5.105 | 5.105
| Std Dev. |  | 0.06364 | 42
| Rel.Std.Dev. |  | 1.247%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 505b (Indicative Value 2.84 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.8 | 1
| 1 | 2 | 2.88 | 7
| Mean |  | 2.84 | 16
| Median |  | 2.84 | 2.84
| Std Dev. |  | 0.056569 | 43
| Rel.Std.Dev. |  | 1.992%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 505b (Indicative Value 1.26 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.26 | 1
| 1 | 2 | 1.26 | 8
| Mean |  | 1.26 | 16
| Median |  | 1.26 | 1.26
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 505b (Indicative Value 18.3 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.1 | 1
| 1 | 2 | 18.5 | 39
| Mean |  | 18.3 | 16
| Median |  | 18.3 | 18.3
| Std Dev. |  | 0.28284 | 45
| Rel.Std.Dev. |  | 1.546%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 505b (Indicative Value 5.98 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.09 | 1
| 1 | 2 | 5.87 | 9
| Mean |  | 5.98 | 16
| Median |  | 5.98 | 5.98
| Std Dev. |  | 0.15556 | 46
| Rel.Std.Dev. |  | 2.601%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 505b (Indicative Value 1.35 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.3 | 1
| 1 | 2 | 1.4 | 4
| Mean |  | 1.35 | 16
| Median |  | 1.35 | 1.35
| Std Dev. |  | 0.070711 | 47
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 505b (Indicative Value 6.2 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.03 | 1
| 1 | 2 | 6.37 | 25
| Mean |  | 6.2 | 16
| Median |  | 6.2 | 6.2
| Std Dev. |  | 0.24042 | 31
| Rel.Std.Dev. |  | 3.878%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 505b (Indicative Value 1.03 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.02 | 1
| 1 | 2 | 1.03 | 10
| Mean |  | 1.025 | 16
| Median |  | 1.025 | 1.025
| Std Dev. |  | 0.0070711 | 32
| Rel.Std.Dev. |  | 0.6899%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 505b (Indicative Value 0.05 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 27
| Mean |  | 0.05 | 16
| Median |  | 0.05 | 0.05
| Std Dev. |  | 0 | 33
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 505b (Indicative Value 35.4 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 35.3 | 1
| 1 | 2 | 35.5 | 28
| Mean |  | 35.4 | 16
| Median |  | 35.4 | 35.4
| Std Dev. |  | 0.14142 | 34
| Rel.Std.Dev. |  | 0.3995%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 505b (Indicative Value 0.36 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.34 | 1
| 1 | 2 | 0.37 | 29
| Mean |  | 0.355 | 16
| Median |  | 0.355 | 0.355
| Std Dev. |  | 0.021213 | 35
| Rel.Std.Dev. |  | 5.976%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 505b (Indicative Value 0.041 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0401 | 1
| 1 | 2 | 0.0417 | 30
| Mean |  | 0.0409 | 16
| Median |  | 0.0409 | 0.0409
| Std Dev. |  | 0.0011314 | 36
| Rel.Std.Dev. |  | 2.766%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 505b (Indicative Value 66 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 66.2 | 1
| 1 | 2 | 66.2 | 31
| Mean |  | 66.2 | 16
| Median |  | 66.2 | 66.2
| Std Dev. |  | 0 | 37
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 505b (Indicative Value 12 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.1 | 1
| 1 | 2 | 11.9 | 6
| Mean |  | 12 | 16
| Median |  | 12 | 12
| Std Dev. |  | 0.14142 | 38
| Rel.Std.Dev. |  | 1.179%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 505b (Indicative Value 32.6 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 33.3 | 1
| 1 | 2 | 31.8 | 7
| Mean |  | 32.55 | 16
| Median |  | 32.55 | 32.55
| Std Dev. |  | 1.0607 | 39
| Rel.Std.Dev. |  | 3.259%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 505b (Indicative Value 48 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46 | 1
| 1 | 2 | 50 | 17
| Mean |  | 48 | 16
| Median |  | 48 | 48
| Std Dev. |  | 2.8284 | 40
| Rel.Std.Dev. |  | 5.893%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 505b (Indicative Value 49.5 ppm) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 49 | 1
| 1 | 2 | 50 | 19
| Mean |  | 49.5 | 16
| Median |  | 49.5 | 49.5
| Std Dev. |  | 0.70711 | 41
| Rel.Std.Dev. |  | 1.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 505b (Indicative Value 8.59 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.6 | 1
| 1 | 2 | 8.58 | 9
| Mean |  | 8.59 | 16
| Median |  | 8.59 | 8.59
| Std Dev. |  | 0.014142 | 42
| Rel.Std.Dev. |  | 0.1646%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 505b (Indicative Value 148 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 148 | 1
| 1 | 2 | 148 | 37
| Mean |  | 148 | 16
| Median |  | 148 | 148
| Std Dev. |  | 0 | 43
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 505b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 38
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 44
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 505b (Indicative Value 10.6 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 10.4 | 1
| 1 | 2 | 10.7 | 39
| Mean |  | 10.55 | 16
| Median |  | 10.55 | 10.55
| Std Dev. |  | 0.21213 | 45
| Rel.Std.Dev. |  | 2.011%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 505b (Indicative Value 8.75 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.7 | 1
| 1 | 2 | 8.8 | 40
| Mean |  | 8.75 | 16
| Median |  | 8.75 | 8.75
| Std Dev. |  | 0.070711 | 46
| Rel.Std.Dev. |  | 0.8081%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 505b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 41
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 47
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 505b (Indicative Value 6.65 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.64 | 1
| 1 | 2 | 6.66 | 7
| Mean |  | 6.65 | 16
| Median |  | 6.65 | 6.65
| Std Dev. |  | 0.014142 | 31
| Rel.Std.Dev. |  | 0.2127%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 505b (Indicative Value 4.3 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.6 | 1
| 1 | 2 | 4 | 14
| Mean |  | 4.3 | 16
| Median |  | 4.3 | 4.3
| Std Dev. |  | 0.42426 | 32
| Rel.Std.Dev. |  | 9.867%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 505b (Indicative Value 262 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 261 | 1
| 1 | 2 | 263 | 15
| Mean |  | 262 | 16
| Median |  | 262 | 262
| Std Dev. |  | 1.4142 | 33
| Rel.Std.Dev. |  | 0.5398%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 505b (Indicative Value 1.04 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.03 | 1
| 1 | 2 | 1.05 | 28
| Mean |  | 1.04 | 16
| Median |  | 1.04 | 1.04
| Std Dev. |  | 0.014142 | 34
| Rel.Std.Dev. |  | 1.36%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 505b (Indicative Value 0.91 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.9 | 1
| 1 | 2 | 0.91 | 29
| Mean |  | 0.905 | 16
| Median |  | 0.905 | 0.905
| Std Dev. |  | 0.0070711 | 35
| Rel.Std.Dev. |  | 0.7813%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 505b (Indicative Value < 0.2 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 0.2 | 1
| 1 | 2 | < 0.2 | 30
| Mean |  |  | 16
| Median |  |  | < 0.2
| Std Dev. |  |  | 36
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 505b (Indicative Value 12.9 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 13 | 1
| 1 | 2 | 12.7 | 31
| Mean |  | 12.85 | 16
| Median |  | 12.85 | 12.85
| Std Dev. |  | 0.21213 | 37
| Rel.Std.Dev. |  | 1.651%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 505b (Indicative Value 0.355 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.351 | 1
| 1 | 2 | 0.358 | 32
| Mean |  | 0.3545 | 16
| Median |  | 0.3545 | 0.3545
| Std Dev. |  | 0.0049497 | 38
| Rel.Std.Dev. |  | 1.396%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 505b (Indicative Value 0.4 ppm) |  |  | Indicative
Tl | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.4 | 33
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.4
| Std Dev. |  | 0 | 39
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 505b (Indicative Value 0.4 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | 0.39 | 8
| Mean |  | 0.395 | 16
| Median |  | 0.395 | 0.395
| Std Dev. |  | 0.0070711 | 40
| Rel.Std.Dev. |  | 1.79%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 505b (Indicative Value 4.19 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.2 | 1
| 1 | 2 | 4.18 | 35
| Mean |  | 4.19 | 16
| Median |  | 4.19 | 4.19
| Std Dev. |  | 0.014142 | 41
| Rel.Std.Dev. |  | 0.3375%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 505b (Indicative Value 85 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 84.1 | 1
| 1 | 2 | 85 | 17
| Mean |  | 84.55 | 16
| Median |  | 84.55 | 84.55
| Std Dev. |  | 0.6364 | 42
| Rel.Std.Dev. |  | 0.7527%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 505b (Indicative Value 9.75 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 10 | 1
| 1 | 2 | 9.5 | 37
| Mean |  | 9.75 | 16
| Median |  | 9.75 | 9.75
| Std Dev. |  | 0.35355 | 43
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 505b (Indicative Value 26.7 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 26.7 | 1
| 1 | 2 | 26.7 | 38
| Mean |  | 26.7 | 16
| Median |  | 26.7 | 26.7
| Std Dev. |  | 0 | 44
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 505b (Indicative Value 2.54 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.52 | 1
| 1 | 2 | 2.55 | 39
| Mean |  | 2.535 | 16
| Median |  | 2.535 | 2.535
| Std Dev. |  | 0.021213 | 45
| Rel.Std.Dev. |  | 0.8368%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 505b (Indicative Value 98 ppm) |  |  | Indicative
Zn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 95 | 1
| 1 | 2 | 100 | 18
| Mean |  | 97.5 | 16
| Median |  | 97.5 | 97.5
| Std Dev. |  | 3.5355 | 46
| Rel.Std.Dev. |  | 3.626%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 505b (Indicative Value 220 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 219 | 1
| 1 | 2 | 221 | 19
| Mean |  | 220 | 16
| Median |  | 220 | 220
| Std Dev. |  | 1.4142 | 47
| Rel.Std.Dev. |  | 0.6428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C713: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.