## Uncertainty & Tolerance Limits
Table 1. Certified Values, Expanded Uncertainty and Tolerance Limits for OREAS 627
Constituent | Certified | Expanded Uncertainty |  | 95% Tolerance Limits
| Value | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 1.8785 | 1.8555 | 1.9014 | 1.8726 | 1.8843
Aqua Regia Digestion (sample weights 10-50g)
Au, Gold (ppm) | 1.8207 | 1.7762 | 1.8653 | 1.8145 | 1.8269
4-Acid Digestion
Ag, Silver (ppm) | 62.7682 | 61.1124 | 64.4241 | 61.7486 | 63.7879
Al, Aluminium (wt.%) | 5.4886 | 5.3455 | 5.6317 | 5.3932 | 5.5841
As, Arsenic (ppm) | 309.018 | 298.9248 | 319.1112 | 302.1206 | 315.9154
Be, Beryllium (ppm) | 1.806 | 1.7221 | 1.89 | 1.7552 | 1.8569
Bi, Bismuth (ppm) | 44.5835 | 42.4125 | 46.7545 | 43.2815 | 45.8855
Ca, Calcium (wt.%) | 0.77627 | 0.74942 | 0.80313 | 0.75546 | 0.79709
Cd, Cadmium (ppm) | 259.3345 | 250.3167 | 268.3522 | 251.1163 | 267.5526
Ce, Cerium (ppm) | 53.2716 | 47.4487 | 59.0945 | 49.9195 | 56.6238
Co, Cobalt (ppm) | 7.5366 | 7.1328 | 7.9403 | 7.3132 | 7.7599
Cr, Chromium (ppm) | 18.1099 | 15.8168 | 20.4031 | 17.081 | 19.1388
Cs, Caesium (ppm) | 4.9344 | 4.7115 | 5.1573 | 4.8091 | 5.0597
Cu, Copper (wt.%) | 0.47701 | 0.46495 | 0.48907 | 0.46933 | 0.4847
Dy, Dysprosium (ppm) | 3.4357 | 3.0963 | 3.7751 | 3.1659 | 3.7055
Er, Erbium (ppm) | 1.4647 | 1.2392 | 1.6903 | 1.3759 | 1.5535
Eu, Europium (ppm) | 1.3242 | 1.1992 | 1.4491 | 1.2376 | 1.4107
Fe, Iron (wt.%) | 6.8599 | 6.6916 | 7.0281 | 6.7515 | 6.9682
Ga, Gallium (ppm) | 18.8694 | 17.9124 | 19.8264 | 18.3999 | 19.3389
Gd, Gadolinium (ppm) | 4.8698 | 3.8949 | 5.8447 | 4.53 | 5.2096
Hf, Hafnium (ppm) | 4.3802 | 4.1388 | 4.6215 | 4.2478 | 4.5126
Ho, Holmium (ppm) | 0.57484 | 0.50754 | 0.64214 | 0.54044 | 0.60923
In, Indium (ppm) | 8.8261 | 8.4416 | 9.2106 | 8.6312 | 9.021
K, Potassium (wt.%) | 2.513 | 2.4432 | 2.5827 | 2.4657 | 2.5602
La, Lanthanum (ppm) | 22.5771 | 18.9311 | 26.223 | 21.4904 | 23.6637
Li, Lithium (ppm) | 17.708 | 16.6179 | 18.7982 | 17.1598 | 18.2563
Lu, Lutetium (ppm) | 0.19132 | 0.16307 | 0.21957 | 0.16652 | 0.21612
Mg, Magnesium (wt.%) | 0.38158 | 0.36556 | 0.3976 | 0.37021 | 0.39295
Mn, Manganese (wt.%) | 0.70342 | 0.68363 | 0.72321 | 0.69113 | 0.71571
Mo, Molybdenum (ppm) | 31.3337 | 30.0714 | 32.596 | 30.6328 | 32.0346
Na, Sodium (wt.%) | 1.0671 | 1.035 | 1.0991 | 1.0463 | 1.0878
Nb, Niobium (ppm) | 10.374 | 9.7089 | 11.0391 | 10.0063 | 10.7417
Nd, Neodymium (ppm) | 31.3774 | 27.0709 | 35.684 | 30.2766 | 32.4783
Ni, Nickel (ppm) | 7.6919 | 7.0306 | 8.3532 | 7.3115 | 8.0723
P, Phosphorus (wt.%) | 0.037017 | 0.035425 | 0.038609 | 0.035809 | 0.038225
Pb, Lead (wt.%) | 2.2646 | 2.2101 | 2.3191 | 2.2289 | 2.3003
Pr, Praseodymium (ppm) | 7.528 | 5.7146 | 9.3414 | 6.8073 | 8.2487
Rb, Rubidium (ppm) | 122.3809 | 115.2088 | 129.5529 | 118.7565 | 126.0052
Re, Rhenium (ppm) | < 0.005 | IND | IND | IND | IND
S, Sulphur (wt.%) | 10.0035 | 9.2999 | 10.7071 | 9.782 | 10.225
Sb, Antimony (ppm) | 142.0611 | 136.6812 | 147.441 | 138.8993 | 145.2229
Sc, Scandium (ppm) | 5.1311 | 4.7703 | 5.4919 | 4.9508 | 5.3114
Sm, Samarium (ppm) | 6.2419 | 5.6596 | 6.8242 | 5.7872 | 6.6966
Sn, Tin (ppm) | 3.4185 | 3.2183 | 3.6188 | 3.2323 | 3.6048
Sr, Strontium (ppm) | 90.5273 | 85.0513 | 96.0033 | 87.8984 | 93.1562
Ta, Tantalum (ppm) | 0.78976 | 0.71981 | 0.85971 | 0.74613 | 0.83339
Tb, Terbium (ppm) | 0.61451 | 0.53311 | 0.69591 | 0.5722 | 0.65682
Te, Tellurium (ppm) | 0.12573 | 0.088186 | 0.16328 | IND | IND
Th, Thorium (ppm) | 9.7251 | 8.8471 | 10.6031 | 9.2476 | 10.2026
Ti, Titanium (wt.%) | 0.13659 | 0.13 | 0.14319 | 0.13279 | 0.1404
Tl, Thallium (ppm) | 14.2142 | 13.5266 | 14.9018 | 13.8054 | 14.6231
Tm, Thulium (ppm) | 0.20606 | 0.18583 | 0.22628 | 0.19307 | 0.21905
U, Uranium (ppm) | 4.8287 | 4.6061 | 5.0514 | 4.6896 | 4.9679
V, Vanadium (ppm) | 23.7633 | 22.5884 | 24.9382 | 22.8894 | 24.6372
W, Tungsten (ppm) | 8.2543 | 7.7961 | 8.7126 | 7.9132 | 8.5954
Y, Yttrium (ppm) | 13.8965 | 12.8665 | 14.9264 | 13.3931 | 14.3998
Yb, Ytterbium (ppm) | 1.3177 | 1.1661 | 1.4692 | 1.2071 | 1.4282
Zn, Zinc (wt.%) | 10.0361 | 9.8192 | 10.253 | 9.8951 | 10.177
Zr, Zirconium (ppm) | 154.0845 | 148.0241 | 160.1449 | 150.7063 | 157.4626
Aqua Regia Digestion
Ag, Silver (ppm) | 62.0419 | 60.507 | 63.5767 | 60.6471 | 63.4366
Al, Aluminium (wt.%) | 0.75513 | 0.70571 | 0.80454 | 0.72763 | 0.78263
As, Arsenic (ppm) | 298.8152 | 289.5975 | 308.0329 | 292.479 | 305.1514
B, Boron (ppm) | < 10 | IND | IND | IND | IND
Be, Beryllium (ppm) | 0.47359 | 0.42942 | 0.51776 | 0.4464 | 0.50077
Bi, Bismuth (ppm) | 44.1561 | 42.2741 | 46.038 | 42.6625 | 45.6497
Ca, Calcium (wt.%) | 0.64541 | 0.61973 | 0.67109 | 0.62173 | 0.66908
Cd, Cadmium (ppm) | 250.742 | 242.4344 | 259.0496 | 245.8349 | 255.649
Ce, Cerium (ppm) | 35.1863 | 31.7541 | 38.6185 | 34.2665 | 36.1061
Co, Cobalt (ppm) | 7.0991 | 6.7082 | 7.4899 | 6.8696 | 7.3285
Cr, Chromium (ppm) | 15.9442 | 14.663 | 17.2253 | 15.0709 | 16.8174
Cs, Caesium (ppm) | 1.6743 | 1.5462 | 1.8023 | 1.5937 | 1.7548
Cu, Copper (wt.%) | 0.4809 | 0.46713 | 0.49468 | 0.47304 | 0.48877
Fe, Iron (wt.%) | 6.3378 | 6.174 | 6.5017 | 6.2018 | 6.4739
Ga, Gallium (ppm) | 3.9743 | 3.6461 | 4.3024 | 3.8061 | 4.1424
Ge, Germanium (ppm) | 0.095381 | 0.065688 | 0.12507 | IND | IND
Hf, Hafnium (ppm) | 1.0681 | 0.9561 | 1.18 | 1.0217 | 1.1144
Hg, Mercury (ppm) | 3.4225 | 3.21 | 3.635 | 3.2607 | 3.5843
In, Indium (ppm) | 8.8014 | 8.3891 | 9.2137 | 8.5245 | 9.0783
K, Potassium (wt.%) | 0.31773 | 0.29667 | 0.3388 | 0.306 | 0.32947
La, Lanthanum (ppm) | 15.6447 | 14.3548 | 16.9345 | 15.0418 | 16.2475
Li, Lithium (ppm) | 5.1819 | 4.4646 | 5.8992 | 4.7886 | 5.5753
Lu, Lutetium (ppm) | 0.063728 | 0.04794 | 0.079516 | IND | IND
Mg, Magnesium (wt.%) | 0.23061 | 0.21793 | 0.24328 | 0.22042 | 0.24079
Mn, Manganese (wt.%) | 0.64421 | 0.62365 | 0.66476 | 0.6301 | 0.65831
Mo, Molybdenum (ppm) | 28.7071 | 27.0928 | 30.3214 | 27.9108 | 29.5034
Na, Sodium (wt.%) | 0.041092 | 0.03829 | 0.043894 | 0.03979 | 0.042394
Nb, Niobium (ppm) | 0.7903 | 0.67876 | 0.90184 | 0.7324 | 0.84821
Ni, Nickel (ppm) | 7.0092 | 6.4434 | 7.575 | 6.6687 | 7.3497
P, Phosphorus (wt.%) | 0.031592 | 0.030161 | 0.033022 | 0.03048 | 0.032703
Pb, Lead (wt.%) | 2.2375 | 2.1934 | 2.2817 | 2.1999 | 2.2752
Rb, Rubidium (ppm) | 17.6513 | 16.3966 | 18.9059 | 17.0083 | 18.2942
S, Sulphur (wt.%) | 9.7098 | 9.2706 | 10.1491 | 9.5122 | 9.9075
Sb, Antimony (ppm) | 112.554 | 106.2373 | 118.8707 | 109.4259 | 115.6821
Sc, Scandium (ppm) | 1.3064 | 1.1438 | 1.4691 | IND | IND
Sn, Tin (ppm) | 1.1392 | 0.99442 | 1.284 | IND | IND
Sr, Strontium (ppm) | 16.6559 | 15.4553 | 17.8565 | 16.1306 | 17.1812
Ta, Tantalum (ppm) | 0.0087778 | 0.0044995 | 0.013056 | IND | IND
Tb, Terbium (ppm) | 0.33996 | 0.30191 | 0.378 | 0.32437 | 0.35554
Te, Tellurium (ppm) | 0.11419 | 0.092969 | 0.13541 | IND | IND
Th, Thorium (ppm) | 6.5392 | 6.1866 | 6.8917 | 6.2892 | 6.7892
Ti, Titanium (wt.%) | 0.028857 | 0.025493 | 0.03222 | 0.027768 | 0.029945
Tl, Thallium (ppm) | 5.2815 | 5.0263 | 5.5368 | 5.0759 | 5.4871
U, Uranium (ppm) | 2.5185 | 2.3628 | 2.6741 | 2.4165 | 2.6205
V, Vanadium (ppm) | 8.4201 | 7.5419 | 9.2984 | 7.7305 | 9.1097
W, Tungsten (ppm) | 4.2946 | 4.0402 | 4.549 | 4.0361 | 4.553
Y, Yttrium (ppm) | 6.2712 | 5.9168 | 6.6257 | 6.0621 | 6.4803
Yb, Ytterbium (ppm) | 0.43899 | 0.36101 | 0.51696 | IND | IND
Zn, Zinc (wt.%) | 10.0724 | 9.8233 | 10.3215 | 9.9134 | 10.2314
Zr, Zirconium (ppm) | 37.0008 | 33.9052 | 40.0964 | 35.8018 | 38.1997
Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.6503 | 5.5185 | 5.782 | 5.5654 | 5.7351
As, Arsenic (ppm) | 315.2944 | 297.8902 | 332.6985 | 307.0653 | 323.5235
Ba, Barium (ppm) | 4383.897 | 4200.2672 | 4567.5267 | 4313.8447 | 4453.9493
Bi, Bismuth (ppm) | 43.6962 | 40.1919 | 47.2004 | 42.1887 | 45.2036
Ca, Calcium (wt.%) | 0.78168 | 0.74126 | 0.8221 | 0.75931 | 0.80405
Cd, Cadmium (ppm) | 262.9101 | 251.3435 | 274.4767 | 257.4101 | 268.4102
Ce, Cerium (ppm) | 70.4741 | 63.8344 | 77.1138 | 67.7739 | 73.1742
Co, Cobalt (ppm) | 8.8723 | 7.7851 | 9.9595 | 7.6723 | 10.0723
Cs, Caesium (ppm) | 5.1434 | 4.5536 | 5.7332 | 4.788 | 5.4988
Cu, Copper (wt.%) | 0.48222 | 0.46946 | 0.49497 | 0.476 | 0.48844
Dy, Dysprosium (ppm) | 3.4207 | 3.079 | 3.7625 | 3.2542 | 3.5873
Er, Erbium (ppm) | 1.6874 | 1.4655 | 1.9093 | 1.6099 | 1.7649
Eu, Europium (ppm) | 1.4216 | 1.114 | 1.7293 | 1.3341 | 1.5092
Fe, Iron (wt.%) | 6.8179 | 6.6621 | 6.9738 | 6.7394 | 6.8964
Ga, Gallium (ppm) | 18.347 | 16.5918 | 20.1022 | 17.674 | 19.02
Gd, Gadolinium (ppm) | 4.9917 | 4.5882 | 5.3952 | 4.7063 | 5.277
Ge, Germanium (ppm) | 2.2228 | 1.8144 | 2.6313 | IND | IND
Ho, Holmium (ppm) | 0.63142 | 0.50832 | 0.75453 | 0.60232 | 0.66053
In, Indium (ppm) | 9.1184 | 8.31 | 9.9268 | 8.7755 | 9.4613
K, Potassium (wt.%) | 2.5635 | 2.4822 | 2.6447 | 2.5235 | 2.6034
La, Lanthanum (ppm) | 34.4574 | 32.4788 | 36.436 | 33.3805 | 35.5342
Li, Lithium (ppm) | 20.0726 | 17.6803 | 22.4648 | 16.7745 | 23.3706
Lu, Lutetium (ppm) | 0.20998 | 0.18998 | 0.22998 | IND | IND
Mg, Magnesium (wt.%) | 0.39593 | 0.37834 | 0.41352 | 0.38332 | 0.40855
Mn, Manganese (wt.%) | 0.71122 | 0.69439 | 0.72804 | 0.7022 | 0.72024
Mo, Molybdenum (ppm) | 32.4602 | 29.8244 | 35.096 | 30.7797 | 34.1407
Nb, Niobium (ppm) | 11.1891 | 9.5506 | 12.8277 | IND | IND
Nd, Neodymium (ppm) | 30.0793 | 27.7248 | 32.4338 | 29.2457 | 30.9128
P, Phosphorus (wt.%) | 0.037643 | 0.027762 | 0.047523 | IND | IND
Pb, Lead (wt.%) | 2.2071 | 2.1527 | 2.2616 | 2.178 | 2.2363
Pr, Praseodymium (ppm) | 8.2696 | 7.6999 | 8.8392 | 7.9902 | 8.5489
Rb, Rubidium (ppm) | 126.9695 | 118.8583 | 135.0808 | 123.5637 | 130.3754
S, Sulphur (wt.%) | 10.7512 | 10.4396 | 11.0627 | 10.553 | 10.9494
Sb, Antimony (ppm) | 147.7445 | 137.0282 | 158.4607 | 144.5918 | 150.8971
Sc, Scandium (ppm) | 5.4947 | 4.3653 | 6.624 | IND | IND
Si, Silicon (wt.%) | 22.4397 | 21.9285 | 22.951 | 22.0627 | 22.8168
Sm, Samarium (ppm) | 5.7198 | 5.2554 | 6.1842 | 5.4692 | 5.9703
Sr, Strontium (ppm) | 113.0918 | 108.9019 | 117.2816 | 110.8494 | 115.3342
Tb, Terbium (ppm) | 0.68256 | 0.60327 | 0.76185 | 0.65154 | 0.71358
Th, Thorium (ppm) | 12.0613 | 10.9989 | 13.1238 | 11.6886 | 12.434
Ti, Titanium (wt.%) | 0.15157 | 0.14511 | 0.15804 | 0.145 | 0.15815
Tl, Thallium (ppm) | 14.4411 | 13.382 | 15.5001 | 13.9981 | 14.8841
Tm, Thulium (ppm) | 0.23297 | 0.15982 | 0.30612 | 0.21278 | 0.25316
U, Uranium (ppm) | 5.2325 | 4.7827 | 5.6823 | 5.021 | 5.444
V, Vanadium (ppm) | 23.0373 | 20.2632 | 25.8115 | 21.6836 | 24.3911
W, Tungsten (ppm) | 8.4497 | 6.9233 | 9.9761 | 8.1924 | 8.707
Y, Yttrium (ppm) | 18.0326 | 16.9765 | 19.0887 | 17.413 | 18.6522
Yb, Ytterbium (ppm) | 1.5036 | 1.3234 | 1.6838 | IND | IND
Zn, Zinc (wt.%) | 10.0404 | 9.8206 | 10.2601 | 9.9497 | 10.1311
Infrared Combustion
S, Sulphur (wt.%) | 11.0498 | 10.8318 | 11.2678 | 10.9404 | 11.1592
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Indicative Values
Table 2. Indicative Values for OREAS 627
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Ag | ppm | 57 | Pd | ppb | 0.66667 | Pt | ppb | < 10
4-Acid Digestion
B | ppm | 36.8333 | Ge | ppm | 0.21774 | Se | ppm | 2.4176
Ba | ppm | 184.7044 | Hg | ppm | 2.591
Aqua Regia Digestion
Ba | ppm | 42.5069 | Ho | ppm | 0.2572 | Re | ppm | 0.0012917
Dy | ppm | 1.5909 | Nd | ppm | 15.6579 | Se | ppm | 2.3602
Er | ppm | 0.61063 | Pd | ppb | 13.1667 | Sm | ppm | 3.0104
Eu | ppm | 0.66558 | Pr | ppm | 3.9915 | Tm | ppm | 0.07256
Gd | ppm | 2.6427 | Pt | ppb | < 5
Peroxide Fusion ICP
Ag | ppm | 63.7625 | Hg | ppm | < 5 | Sn | ppm | 6.9382
B | ppm | 26.7556 | Na | wt.% | 1.0817 | Ta | ppm | 0.85322
Be | ppm | 1.7838 | Ni | ppm | 17.0402 | Te | ppm | < 5
Cr | ppm | 55.0184 | Re | ppm | < 0.05 | Zr | ppm | 181.8965
Hf | ppm | 4.456 | Se | ppm | < 5
Infrared Combustion
C | wt.% | 0.34167
Borate Fusion XRF
Al2O3 | wt.% | 10.7751 | MgO | wt.% | 0.66209 | Sn | ppm | 3.7702
As | ppm | 9071.4902 | Mn | wt.% | 0.72498 | Sr | ppm | 71.237
Ba | ppm | 4568.0553 | Na2O | wt.% | 1.605 | TiO2 | wt.% | 0.24599
CaO | wt.% | 1.0943 | Ni | ppm | 9.7641 | V | ppm | 19.2443
Cr | ppm | 35.435 | P | wt.% | 0.037408 | Zn | wt.% | 10.5507
Cu | wt.% | 0.4431 | Pb | wt.% | 2.3665 | Zr | ppm | 171.2761
Fe | wt.% | 6.9653 | S | wt.% | 7.3282
K2O | wt.% | 3.0973 | SiO2 | wt.% | 48.0621
Thermogravimetry
LOI1000 | wt.% | 8.29
Laser Ablation ICP-MS
Ag | ppm | 57.45 | Hf | ppm | 4.91 | Se | ppm | < 5
As | ppm | 319 | Ho | ppm | 0.64 | Sm | ppm | 5.855
Ba | ppm | 4305 | In | ppm | 8.125 | Sn | ppm | 6.4
Be | ppm | 1.8 | La | ppm | 34.3 | Sr | ppm | 102
Bi | ppm | 45.7 | Lu | ppm | 0.205 | Ta | ppm | 0.81
Ce | ppm | 66.4 | Mn | wt.% | 0.7235 | Tb | ppm | 0.66
Co | ppm | 7.85 | Mo | ppm | 30.4 | Te | ppm | 0.25
Cr | ppm | 18 | Nb | ppm | 11.1 | Th | ppm | 11.75
Cs | ppm | 4.97 | Nd | ppm | 30.85 | Ti | wt.% | 0.1555
Cu | wt.% | 0.5145 | Ni | ppm | 15 | Tm | ppm | 0.22
Dy | ppm | 3.565 | Pb | wt.% | 2.09 | U | ppm | 4.72
Er | ppm | 1.565 | Pr | ppm | 8.31 | V | ppm | 23
Eu | ppm | 1.315 | Rb | ppm | 124.5 | W | ppm | 8.25
Ga | ppm | 17.85 | Re | ppm | < 0.01 | Y | ppm | 18
Gd | ppm | 4.615 | Sb | ppm | 152 | Yb | ppm | 1.525
Ge | ppm | 2.05 | Sc | ppm | 5.7 | Zr | ppm | 173
SI unit equivalents: ppb (parts per billion; 1 x 10⁹) ≡ µg/kg; ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Performance Gates
Table 3. Certified Values and Performance Gates for OREAS 627
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 1.8785 | 0.063209 | 1.7521 | 2.0049 | 1.6889 | 2.0681 | 3.365% | 6.73% | 10.09% | 1.7846 | 1.9724
Aqua Regia Digestion (sample weights 10-50g)
Au, ppm | 1.8207 | 0.091997 | 1.6367 | 2.0047 | 1.5447 | 2.0967 | 5.053% | 10.11% | 15.16% | 1.7297 | 1.9118
4-Acid Digestion
Ag, ppm | 62.7682 | 1.7977 | 59.1728 | 66.3637 | 57.3751 | 68.1614 | 2.864% | 5.728% | 8.592% | 59.6298 | 65.9067
Al, wt.% | 5.4886 | 0.17637 | 5.1359 | 5.8414 | 4.9595 | 6.0177 | 3.213% | 6.427% | 9.64% | 5.2142 | 5.763
As, ppm | 309.018 | 13.8328 | 281.3524 | 336.6836 | 267.5196 | 350.5164 | 4.476% | 8.953% | 13.43% | 293.5671 | 324.4689
Be, ppm | 1.806 | 0.10611 | 1.5938 | 2.0182 | 1.4877 | 2.1244 | 5.876% | 11.75% | 17.63% | 1.7157 | 1.8963
Bi, ppm | 44.5835 | 2.8671 | 38.8493 | 50.3177 | 35.9821 | 53.1848 | 6.431% | 12.86% | 19.29% | 42.3543 | 46.8126
Ca, wt.% | 0.77627 | 0.037111 | 0.70205 | 0.8505 | 0.66494 | 0.88761 | 4.781% | 9.561% | 14.34% | 0.73746 | 0.81509
Cd, ppm | 259.3345 | 11.9339 | 235.4667 | 283.2022 | 223.5328 | 295.1361 | 4.602% | 9.203% | 13.81% | 246.3677 | 272.3012
Ce, ppm | 53.2716 | 10.1399 | 32.9918 | 73.5514 | 22.8519 | 83.6913 | 19.03% | 38.07% | 57.1% | 50.6081 | 55.9352
Co, ppm | 7.5366 | 0.56579 | 6.405 | 8.6682 | 5.8392 | 9.234 | 7.507% | 15.01% | 22.52% | 7.1597 | 7.9134
Cr, ppm | 18.1099 | 2.7104 | 12.6891 | 23.5307 | 9.9787 | 26.2411 | 14.97% | 29.93% | 44.9% | 17.2044 | 19.0154
Cs, ppm | 4.9344 | 0.22688 | 4.4806 | 5.3882 | 4.2537 | 5.615 | 4.598% | 9.196% | 13.79% | 4.6877 | 5.1811
Cu, wt.% | 0.47701 | 0.019705 | 0.4376 | 0.51642 | 0.4179 | 0.53613 | 4.131% | 8.262% | 12.39% | 0.45316 | 0.50086
Dy, ppm | 3.4357 | 0.19853 | 3.0386 | 3.8327 | 2.8401 | 4.0313 | 5.778% | 11.56% | 17.34% | 3.2639 | 3.6075
Er, ppm | 1.4647 | 0.16629 | 1.1322 | 1.7973 | 0.96587 | 1.9636 | 11.35% | 22.71% | 34.06% | 1.3915 | 1.538
Eu, ppm | 1.3242 | 0.047358 | 1.2295 | 1.4189 | 1.1821 | 1.4662 | 3.576% | 7.153% | 10.73% | 1.258 | 1.3904
Fe, wt.% | 6.8599 | 0.23027 | 6.3993 | 7.3204 | 6.1691 | 7.5507 | 3.357% | 6.713% | 10.07% | 6.5169 | 7.2029
Ga, ppm | 18.8694 | 1.1045 | 16.6603 | 21.0785 | 15.5558 | 22.183 | 5.854% | 11.71% | 17.56% | 17.9259 | 19.8129
Gd, ppm | 4.8698 | 0.8271 | 3.2156 | 6.524 | 2.3885 | 7.3511 | 16.98% | 33.97% | 50.95% | 4.6263 | 5.1133
Hf, ppm | 4.3802 | 0.21445 | 3.9513 | 4.8091 | 3.7368 | 5.0235 | 4.896% | 9.792% | 14.69% | 4.1612 | 4.5992
Ho, ppm | 0.57484 | 0.034815 | 0.5052 | 0.64447 | 0.47039 | 0.67928 | 6.057% | 12.11% | 18.17% | 0.54609 | 0.60358
In, ppm | 8.8261 | 0.44832 | 7.9294 | 9.7227 | 7.4811 | 10.171 | 5.079% | 10.16% | 15.24% | 8.3848 | 9.2674
K, wt.% | 2.513 | 0.10659 | 2.2998 | 2.7261 | 2.1932 | 2.8327 | 4.241% | 8.483% | 12.72% | 2.3873 | 2.6386
La, ppm | 22.5771 | 7.2714 | 8.0343 | 37.1198 | 0.76294 | 44.3912 | 32.21% | 64.41% | 96.62% | 21.4482 | 23.7059
Li, ppm | 17.708 | 1.1527 | 15.4027 | 20.0134 | 14.25 | 21.1661 | 6.509% | 13.02% | 19.53% | 16.8226 | 18.5934
Lu, ppm | 0.19132 | 0.023616 | 0.14409 | 0.23855 | 0.12047 | 0.26216 | 12.34% | 24.69% | 37.03% | 0.18175 | 0.20088
Mg, wt.% | 0.38158 | 0.023139 | 0.3353 | 0.42786 | 0.31216 | 0.45099 | 6.064% | 12.13% | 18.19% | 0.3625 | 0.40066
Mn, wt.% | 0.70342 | 0.030802 | 0.64182 | 0.76503 | 0.61102 | 0.79583 | 4.379% | 8.758% | 13.14% | 0.66825 | 0.73859
Mo, ppm | 31.3337 | 1.6014 | 28.1309 | 34.5366 | 26.5294 | 36.138 | 5.111% | 10.22% | 15.33% | 29.767 | 32.9004
Na, wt.% | 1.0671 | 0.036694 | 0.99368 | 1.1405 | 0.95698 | 1.1771 | 3.439% | 6.878% | 10.32% | 1.0137 | 1.1204
Nb, ppm | 10.374 | 1.0111 | 8.3518 | 12.3962 | 7.3407 | 13.4073 | 9.747% | 19.49% | 29.24% | 9.8553 | 10.8927
Nd, ppm | 31.3774 | 2.8784 | 25.6206 | 37.1342 | 22.7422 | 40.0126 | 9.173% | 18.35% | 27.52% | 29.8085 | 32.9463
Ni, ppm | 7.6919 | 0.50849 | 6.6749 | 8.7089 | 6.1664 | 9.2174 | 6.611% | 13.22% | 19.83% | 7.3073 | 8.0765
P, wt.% | 0.037017 | 0.0019188 | 0.033179 | 0.040855 | 0.031261 | 0.042773 | 5.184% | 10.37% | 15.55% | 0.035166 | 0.038868
Pb, wt.% | 2.2646 | 0.066866 | 2.1309 | 2.3983 | 2.064 | 2.4652 | 2.953% | 5.905% | 8.858% | 2.1514 | 2.3778
Pr, ppm | 7.528 | 1.553 | 4.422 | 10.634 | 2.869 | 12.187 | 20.63% | 41.26% | 61.89% | 7.1516 | 7.9044
Rb, ppm | 122.3809 | 9.6277 | 103.1254 | 141.6363 | 93.4977 | 151.2641 | 7.867% | 15.73% | 23.6% | 116.2618 | 128.4999
Re, ppm | < 0.005 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, wt.% | 10.0035 | 0.98244 | 8.0386 | 11.9684 | 7.0562 | 12.9508 | 9.821% | 19.64% | 29.46% | 9.5033 | 10.5037
Sb, ppm | 142.0611 | 6.2254 | 129.6104 | 154.5118 | 123.3851 | 160.7372 | 4.382% | 8.764% | 13.15% | 134.9581 | 149.1642
Sc, ppm | 5.1311 | 0.48435 | 4.1624 | 6.0998 | 3.6781 | 6.5842 | 9.439% | 18.88% | 28.32% | 4.8746 | 5.3877
Sm, ppm | 6.2419 | 0.30361 | 5.6347 | 6.8491 | 5.3311 | 7.1527 | 4.864% | 9.728% | 14.59% | 5.9298 | 6.554
Sn, ppm | 3.4185 | 0.1856 | 3.0473 | 3.7897 | 2.8617 | 3.9753 | 5.429% | 10.86% | 16.29% | 3.2476 | 3.5895
Sr, ppm | 90.5273 | 9.8804 | 70.7666 | 110.2881 | 60.8862 | 120.1685 | 10.91% | 21.83% | 32.74% | 86.001 | 95.0537
Ta, ppm | 0.78976 | 0.1083 | 0.57316 | 1.0064 | 0.46486 | 1.1147 | 13.71% | 27.43% | 41.14% | 0.75027 | 0.82925
Tb, ppm | 0.61451 | 0.087258 | 0.43999 | 0.78903 | 0.35273 | 0.87629 | 14.2% | 28.4% | 42.6% | 0.58378 | 0.64524
Te, ppm | 0.12573 | 0.025387 | 0.074959 | 0.17651 | 0.049572 | 0.20189 | 20.19% | 40.38% | 60.57% | 0.11945 | 0.13202
Th, ppm | 9.7251 | 1.6318 | 6.4616 | 12.9886 | 4.8298 | 14.6204 | 16.78% | 33.56% | 50.34% | 9.2389 | 10.2114
Ti, wt.% | 0.13659 | 0.0091627 | 0.11827 | 0.15492 | 0.10911 | 0.16408 | 6.708% | 13.42% | 20.12% | 0.12976 | 0.14342
Tl, ppm | 14.2142 | 0.81503 | 12.5842 | 15.8443 | 11.7691 | 16.6593 | 5.734% | 11.47% | 17.2% | 13.5035 | 14.9249
Tm, ppm | 0.20606 | 0.0091606 | 0.18774 | 0.22438 | 0.17858 | 0.23354 | 4.446% | 8.891% | 13.34% | 0.19576 | 0.21636
U, ppm | 4.8287 | 0.28413 | 4.2605 | 5.397 | 3.9764 | 5.6811 | 5.884% | 11.77% | 17.65% | 4.5873 | 5.0702
V, ppm | 23.7633 | 1.3466 | 21.0701 | 26.4565 | 19.7236 | 27.803 | 5.667% | 11.33% | 17% | 22.5751 | 24.9515
W, ppm | 8.2543 | 0.5761 | 7.1021 | 9.4065 | 6.526 | 9.9826 | 6.979% | 13.96% | 20.94% | 7.8416 | 8.667
Y, ppm | 13.8965 | 1.767 | 10.3624 | 17.4305 | 8.5953 | 19.1976 | 12.72% | 25.43% | 38.15% | 13.2016 | 14.5913
Yb, ppm | 1.3177 | 0.14231 | 1.033 | 1.6023 | 0.89074 | 1.7446 | 10.8% | 21.6% | 32.4% | 1.2518 | 1.3835
Zn, wt.% | 10.0361 | 0.31247 | 9.4111 | 10.661 | 9.0987 | 10.9735 | 3.113% | 6.227% | 9.34% | 9.5343 | 10.5379
Zr, ppm | 154.0845 | 9.2949 | 135.4948 | 172.6742 | 126.1999 | 181.9691 | 6.032% | 12.06% | 18.1% | 146.3803 | 161.7887
Aqua Regia Digestion
Ag, ppm | 62.0419 | 2.4695 | 57.1028 | 66.9809 | 54.6333 | 69.4504 | 3.98% | 7.961% | 11.94% | 58.9398 | 65.1439
Al, wt.% | 0.75513 | 0.054966 | 0.6452 | 0.86506 | 0.59023 | 0.92003 | 7.279% | 14.56% | 21.84% | 0.71737 | 0.79288
As, ppm | 298.8152 | 13.2097 | 272.3958 | 325.2346 | 259.1861 | 338.4444 | 4.421% | 8.841% | 13.26% | 283.8745 | 313.756
B, ppm | < 10 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Be, ppm | 0.47359 | 0.06527 | 0.34305 | 0.60413 | 0.27778 | 0.6694 | 13.78% | 27.56% | 41.35% | 0.44991 | 0.49727
Bi, ppm | 44.1561 | 2.3227 | 39.5106 | 48.8015 | 37.1879 | 51.1243 | 5.26% | 10.52% | 15.78% | 41.9483 | 46.3639
Ca, wt.% | 0.64541 | 0.022287 | 0.60083 | 0.68998 | 0.57855 | 0.71227 | 3.453% | 6.906% | 10.36% | 0.61314 | 0.67768
Cd, ppm | 250.742 | 12.8079 | 225.1261 | 276.3578 | 212.3182 | 289.1658 | 5.108% | 10.22% | 15.32% | 238.2049 | 263.2791
Ce, ppm | 35.1863 | 5.6196 | 23.9471 | 46.4255 | 18.3275 | 52.0451 | 15.97% | 31.94% | 47.91% | 33.427 | 36.9456
Co, ppm | 7.0991 | 0.62193 | 5.8552 | 8.3429 | 5.2333 | 8.9648 | 8.761% | 17.52% | 26.28% | 6.7441 | 7.454
Cr, ppm | 15.9442 | 1.2066 | 13.5309 | 18.3574 | 12.3243 | 19.564 | 7.568% | 15.14% | 22.7% | 15.1469 | 16.7414
Cs, ppm | 1.6743 | 0.17993 | 1.3144 | 2.0341 | 1.1345 | 2.2141 | 10.75% | 21.49% | 32.24% | 1.5906 | 1.758
Cu, wt.% | 0.4809 | 0.021665 | 0.43757 | 0.52423 | 0.41591 | 0.5459 | 4.505% | 9.01% | 13.52% | 0.45686 | 0.50495
Fe, wt.% | 6.3378 | 0.26511 | 5.8076 | 6.8681 | 5.5425 | 7.1332 | 4.183% | 8.366% | 12.55% | 6.0209 | 6.6547
Ga, ppm | 3.9743 | 0.4564 | 3.0615 | 4.8871 | 2.6051 | 5.3435 | 11.48% | 22.97% | 34.45% | 3.7756 | 4.173
Ge, ppm | 0.095381 | 0.028174 | 0.039033 | 0.15173 | 0.010859 | 0.1799 | 29.54% | 59.08% | 88.62% | 0.090612 | 0.10015
Hf, ppm | 1.0681 | 0.18159 | 0.70488 | 1.4312 | 0.52328 | 1.6128 | 17% | 34% | 51.01% | 1.0147 | 1.1215
Hg, ppm | 3.4225 | 0.27029 | 2.8819 | 3.9631 | 2.6116 | 4.2334 | 7.897% | 15.79% | 23.69% | 3.2514 | 3.5936
In, ppm | 8.8014 | 0.62049 | 7.5604 | 10.0423 | 6.9399 | 10.6628 | 7.05% | 14.1% | 21.15% | 8.3613 | 9.2414
K, wt.% | 0.31773 | 0.039862 | 0.23801 | 0.39746 | 0.19815 | 0.43732 | 12.55% | 25.09% | 37.64% | 0.30185 | 0.33362
La, ppm | 15.6447 | 1.751 | 12.1428 | 19.1466 | 10.3918 | 20.8975 | 11.19% | 22.38% | 33.58% | 14.8624 | 16.4269
Li, ppm | 5.1819 | 0.62439 | 3.9332 | 6.4307 | 3.3088 | 7.0551 | 12.05% | 24.1% | 36.15% | 4.9228 | 5.441
Lu, ppm | 0.063728 | 0.0093947 | 0.044938 | 0.082517 | 0.035544 | 0.091912 | 14.74% | 29.48% | 44.23% | 0.060541 | 0.066914
Mg, wt.% | 0.23061 | 0.014267 | 0.20207 | 0.25914 | 0.1878 | 0.27341 | 6.187% | 12.37% | 18.56% | 0.21908 | 0.24214
Mn, wt.% | 0.64421 | 0.033087 | 0.57803 | 0.71038 | 0.54495 | 0.74347 | 5.136% | 10.27% | 15.41% | 0.612 | 0.67642
Mo, ppm | 28.7071 | 1.8729 | 24.9612 | 32.453 | 23.0883 | 34.3259 | 6.524% | 13.05% | 19.57% | 27.2717 | 30.1425
Na, wt.% | 0.041092 | 0.0069875 | 0.027117 | 0.055067 | 0.020129 | 0.062055 | 17% | 34.01% | 51.01% | 0.039037 | 0.043147
Nb, ppm | 0.7903 | 0.16437 | 0.46155 | 1.1191 | 0.29718 | 1.2834 | 20.8% | 41.6% | 62.4% | 0.75079 | 0.82982
Ni, ppm | 7.0092 | 0.77996 | 5.4493 | 8.5691 | 4.6693 | 9.3491 | 11.13% | 22.26% | 33.38% | 6.6587 | 7.3596
P, wt.% | 0.031592 | 0.0016312 | 0.028329 | 0.034854 | 0.026698 | 0.036485 | 5.163% | 10.33% | 15.49% | 0.030012 | 0.033171
Pb, wt.% | 2.2375 | 0.06644 | 2.1047 | 2.3704 | 2.0382 | 2.4369 | 2.969% | 5.939% | 8.908% | 2.1257 | 2.3494
Rb, ppm | 17.6513 | 1.8514 | 13.9485 | 21.354 | 12.0972 | 23.2054 | 10.49% | 20.98% | 31.47% | 16.7687 | 18.5338
S, wt.% | 9.7098 | 0.48524 | 8.7394 | 10.6803 | 8.2541 | 11.1656 | 4.997% | 9.995% | 14.99% | 9.2243 | 10.1953
Sb, ppm | 112.554 | 10.8876 | 90.7789 | 134.3291 | 79.8913 | 145.2167 | 9.673% | 19.35% | 29.02% | 106.9263 | 118.1817
Sc, ppm | 1.3064 | 0.26695 | 0.77254 | 1.8403 | 0.50559 | 2.1073 | 20.43% | 40.87% | 61.3% | 1.2411 | 1.3718
Sn, ppm | 1.1392 | 0.2065 | 0.7262 | 1.5522 | 0.51971 | 1.7587 | 18.13% | 36.25% | 54.38% | 1.0822 | 1.1962
Sr, ppm | 16.6559 | 2.0634 | 12.529 | 20.7828 | 10.4656 | 22.8462 | 12.39% | 24.78% | 37.17% | 15.8231 | 17.4887
Ta, ppm | 0.0087778 | 0.0021302 | 0.0045173 | 0.013038 | 0.0023871 | 0.015168 | 24.27% | 48.54% | 72.8% | 0.0083389 | 0.0092167
Tb, ppm | 0.33996 | 0.017259 | 0.30544 | 0.37447 | 0.28818 | 0.39173 | 5.077% | 10.15% | 15.23% | 0.32296 | 0.35695
Te, ppm | 0.11419 | 0.016574 | 0.081038 | 0.14734 | 0.064464 | 0.16391 | 14.52% | 29.03% | 43.55% | 0.10848 | 0.1199
Th, ppm | 6.5392 | 0.35842 | 5.8223 | 7.256 | 5.4639 | 7.6144 | 5.481% | 10.96% | 16.44% | 6.2122 | 6.8661
Ti, wt.% | 0.028857 | 0.0061411 | 0.016574 | 0.041139 | 0.010433 | 0.04728 | 21.28% | 42.56% | 63.84% | 0.027414 | 0.030299
Tl, ppm | 5.2815 | 0.24972 | 4.7821 | 5.7809 | 4.5324 | 6.0307 | 4.728% | 9.456% | 14.18% | 5.0174 | 5.5456
U, ppm | 2.5185 | 0.25108 | 2.0163 | 3.0206 | 1.7653 | 3.2717 | 9.969% | 19.94% | 29.91% | 2.3926 | 2.6444
V, ppm | 8.4201 | 0.80972 | 6.8007 | 10.0396 | 5.991 | 10.8493 | 9.616% | 19.23% | 28.85% | 7.9991 | 8.8411
W, ppm | 4.2946 | 0.35692 | 3.5807 | 5.0084 | 3.2238 | 5.3653 | 8.311% | 16.62% | 24.93% | 4.0798 | 4.5093
Y, ppm | 6.2712 | 0.56977 | 5.1317 | 7.4108 | 4.5619 | 7.9805 | 9.086% | 18.17% | 27.26% | 5.9576 | 6.5848
Yb, ppm | 0.43899 | 0.048949 | 0.34109 | 0.53689 | 0.29214 | 0.58584 | 11.15% | 22.3% | 33.45% | 0.41704 | 0.46094
Zn, wt.% | 10.0724 | 0.44362 | 9.1852 | 10.9597 | 8.7416 | 11.4033 | 4.404% | 8.809% | 13.21% | 9.5688 | 10.576
Zr, ppm | 37.0008 | 5.6455 | 25.7097 | 48.2918 | 20.0642 | 53.9373 | 15.26% | 30.52% | 45.77% | 35.1507 | 38.8508
Peroxide Fusion ICP
Al, wt.% | 5.6503 | 0.15539 | 5.3395 | 5.9611 | 5.1841 | 6.1164 | 2.75% | 5.5% | 8.25% | 5.3678 | 5.9328
As, ppm | 315.2944 | 19.1139 | 277.0666 | 353.5222 | 257.9527 | 372.636 | 6.062% | 12.12% | 18.19% | 299.5297 | 331.0591
Ba, ppm | 4383.897 | 258.6282 | 3866.6405 | 4901.1534 | 3608.0123 | 5159.7817 | 5.9% | 11.8% | 17.7% | 4164.7021 | 4603.0918
Bi, ppm | 43.6962 | 3.2803 | 37.1356 | 50.2567 | 33.8554 | 53.537 | 7.507% | 15.01% | 22.52% | 41.5114 | 45.881
Ca, wt.% | 0.78168 | 0.031921 | 0.71784 | 0.84552 | 0.68592 | 0.87745 | 4.084% | 8.167% | 12.25% | 0.7426 | 0.82077
Cd, ppm | 262.9101 | 8.7585 | 245.3931 | 280.4272 | 236.6346 | 289.1857 | 3.331% | 6.663% | 9.994% | 249.7646 | 276.0556
Ce, ppm | 70.4741 | 4.3364 | 61.8012 | 79.1469 | 57.4648 | 83.4834 | 6.153% | 12.31% | 18.46% | 66.9504 | 73.9978
Co, ppm | 8.8723 | 1.1985 | 6.4752 | 11.2694 | 5.2767 | 12.4679 | 13.51% | 27.02% | 40.53% | 8.4287 | 9.3159
Cs, ppm | 5.1434 | 0.23274 | 4.6779 | 5.6089 | 4.4452 | 5.8416 | 4.525% | 9.05% | 13.58% | 4.8862 | 5.4005
Cu, wt.% | 0.48222 | 0.016325 | 0.44957 | 0.51487 | 0.43324 | 0.53119 | 3.385% | 6.771% | 10.16% | 0.45811 | 0.50633
Dy, ppm | 3.4207 | 0.2503 | 2.9202 | 3.9213 | 2.6699 | 4.1716 | 7.317% | 14.63% | 21.95% | 3.2497 | 3.5918
Er, ppm | 1.6874 | 0.119 | 1.4494 | 1.9254 | 1.3304 | 2.0444 | 7.052% | 14.1% | 21.16% | 1.603 | 1.7718
Eu, ppm | 1.4216 | 0.21817 | 0.9853 | 1.858 | 0.76713 | 2.0762 | 15.35% | 30.69% | 46.04% | 1.3506 | 1.4927
Fe, wt.% | 6.8179 | 0.1789 | 6.4601 | 7.1757 | 6.2812 | 7.3546 | 2.624% | 5.248% | 7.872% | 6.477 | 7.1588
Ga, ppm | 18.347 | 1.6327 | 15.0816 | 21.6124 | 13.4489 | 23.2452 | 8.899% | 17.8% | 26.7% | 17.4297 | 19.2644
Gd, ppm | 4.9917 | 0.31024 | 4.3712 | 5.6122 | 4.061 | 5.9224 | 6.215% | 12.43% | 18.65% | 4.7421 | 5.2413
Ge, ppm | 2.2228 | 0.34824 | 1.5263 | 2.9193 | 1.1781 | 3.2676 | 15.67% | 31.33% | 47% | 2.1117 | 2.334
Ho, ppm | 0.63142 | 0.061179 | 0.50907 | 0.75378 | 0.44789 | 0.81496 | 9.689% | 19.38% | 29.07% | 0.59985 | 0.663
In, ppm | 9.1184 | 0.65235 | 7.8137 | 10.4231 | 7.1614 | 11.0755 | 7.154% | 14.31% | 21.46% | 8.6625 | 9.5743
K, wt.% | 2.5635 | 0.051605 | 2.4603 | 2.6667 | 2.4087 | 2.7183 | 2.013% | 4.026% | 6.039% | 2.4353 | 2.6916
La, ppm | 34.4574 | 1.4822 | 31.493 | 37.4218 | 30.0108 | 38.904 | 4.302% | 8.603% | 12.9% | 32.7345 | 36.1803
Li, ppm | 20.0726 | 3.3015 | 13.4695 | 26.6756 | 10.1679 | 29.9772 | 16.45% | 32.9% | 49.34% | 19.0689 | 21.0762
Lu, ppm | 0.20998 | 0.014535 | 0.18091 | 0.23905 | 0.16637 | 0.25358 | 6.922% | 13.84% | 20.77% | 0.19948 | 0.22048
Mg, wt.% | 0.39593 | 0.021195 | 0.35355 | 0.43832 | 0.33235 | 0.45952 | 5.353% | 10.71% | 16.06% | 0.37614 | 0.41573
Mn, wt.% | 0.71122 | 0.019616 | 0.67199 | 0.75045 | 0.65237 | 0.77007 | 2.758% | 5.516% | 8.274% | 0.67566 | 0.74678
Mo, ppm | 32.4602 | 2.8439 | 26.7724 | 38.148 | 23.9284 | 40.992 | 8.761% | 17.52% | 26.28% | 30.8372 | 34.0832
Nb, ppm | 11.1891 | 1.1471 | 8.895 | 13.4833 | 7.7479 | 14.6304 | 10.25% | 20.5% | 30.76% | 10.6297 | 11.7486
Nd, ppm | 30.0793 | 1.9106 | 26.2582 | 33.9004 | 24.3476 | 35.811 | 6.352% | 12.7% | 19.06% | 28.5753 | 31.5833
P, wt.% | 0.037643 | 0.0081433 | 0.021356 | 0.053929 | 0.013213 | 0.062072 | 21.63% | 43.27% | 64.9% | 0.03576 | 0.039525
Pb, wt.% | 2.2071 | 0.076333 | 2.0545 | 2.3598 | 1.9781 | 2.4361 | 3.458% | 6.917% | 10.38% | 2.0968 | 2.3175
Pr, ppm | 8.2696 | 0.29376 | 7.682 | 8.8571 | 7.3883 | 9.1509 | 3.552% | 7.105% | 10.66% | 7.8561 | 8.683
Rb, ppm | 126.9695 | 7.4622 | 112.0452 | 141.8938 | 104.583 | 149.356 | 5.877% | 11.75% | 17.63% | 120.6211 | 133.318
S, wt.% | 10.7512 | 0.32878 | 10.0936 | 11.4087 | 9.7648 | 11.7375 | 3.058% | 6.116% | 9.174% | 10.2136 | 11.2887
Sb, ppm | 147.7445 | 17.9766 | 111.7913 | 183.6976 | 93.8147 | 201.6742 | 12.17% | 24.33% | 36.5% | 140.3572 | 155.1317
Sc, ppm | 5.4947 | 0.61978 | 4.2551 | 6.7342 | 3.6353 | 7.354 | 11.28% | 22.56% | 33.84% | 5.2199 | 5.7694
Si, wt.% | 22.4397 | 0.3534 | 21.7329 | 23.1465 | 21.3795 | 23.4999 | 1.575% | 3.15% | 4.725% | 21.3177 | 23.5617
Sm, ppm | 5.7198 | 0.39959 | 4.9206 | 6.519 | 4.521 | 6.9185 | 6.986% | 13.97% | 20.96% | 5.4338 | 6.0058
Sr, ppm | 113.0918 | 5.8351 | 101.4216 | 124.762 | 95.5865 | 130.5971 | 5.16% | 10.32% | 15.48% | 107.4372 | 118.7464
Tb, ppm | 0.68256 | 0.050775 | 0.58101 | 0.78411 | 0.53023 | 0.83489 | 7.439% | 14.88% | 22.32% | 0.64843 | 0.71669
Th, ppm | 12.0613 | 1.0701 | 9.9211 | 14.2015 | 8.851 | 15.2717 | 8.872% | 17.74% | 26.62% | 11.4583 | 12.6644
Ti, wt.% | 0.15157 | 0.0045205 | 0.14253 | 0.16061 | 0.13801 | 0.16513 | 2.982% | 5.965% | 8.947% | 0.14399 | 0.15915
Tl, ppm | 14.4411 | 0.94229 | 12.5565 | 16.3256 | 11.6142 | 17.2679 | 6.525% | 13.05% | 19.58% | 13.719 | 15.1631
Tm, ppm | 0.23297 | 0.030235 | 0.1725 | 0.29344 | 0.14226 | 0.32367 | 12.98% | 25.96% | 38.93% | 0.22132 | 0.24462
U, ppm | 5.2325 | 0.4497 | 4.3332 | 6.1319 | 3.8835 | 6.5816 | 8.594% | 17.19% | 25.78% | 4.9709 | 5.4942
V, ppm | 23.0373 | 1.96 | 19.1173 | 26.9574 | 17.1572 | 28.9174 | 8.508% | 17.02% | 25.52% | 21.8855 | 24.1892
W, ppm | 8.4497 | 1.1298 | 6.19 | 10.7094 | 5.0602 | 11.8393 | 13.37% | 26.74% | 40.11% | 8.0273 | 8.8722
Y, ppm | 18.0326 | 1.1298 | 15.7729 | 20.2923 | 14.6431 | 21.4221 | 6.266% | 12.53% | 18.8% | 17.131 | 18.9342
Yb, ppm | 1.5036 | 0.10542 | 1.2927 | 1.7144 | 1.1873 | 1.8199 | 7.011% | 14.02% | 21.03% | 1.4284 | 1.5788
Zn, wt.% | 10.0404 | 0.22563 | 9.5891 | 10.4916 | 9.3635 | 10.7172 | 2.247% | 4.494% | 6.742% | 9.5383 | 10.5424
Infrared Combustion
S, wt.% | 11.0498 | 0.31382 | 10.4221 | 11.6774 | 10.1083 | 11.9913 | 2.84% | 5.68% | 8.52% | 10.4973 | 11.6023
SI unit equivalents: ppm (parts per million; 1 x 10⁶) ≡ mg/kg; wt.% (weight per cent) ≡ % (mass fraction)

## Abbreviations
Table 4. Abbreviations used for OREAS 627
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with atomic absorption spectroscopy
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma mass spectroscopy
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digestion with inductively coupled plasma optical emission spectroscopy
4A*OES/AAS | 4-acid (HF-HNO3-HClO4-HCl) digestion with AAS or ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digestion with ICP-OES or ICP-MS finish
ABL*MS | laser ablation with inductively coupled plasma mass spectroscopy
AR*AAS | aqua regia digestion with atomic absorption spectroscopy
AR*MS | aqua regia digestion with inductively coupled plasma mass spectroscopy
AR*OES | aqua regia digestion with inductively coupled plasma optical emission spectroscopy
AR*OES/AAS | aqua regia digestion with ICP-OES or AAS finish
AR*OES/MS | aqua regia digestion with ICP-OES or ICP-MS finish
BF*MS | lithium borate fusion with inductively coupled plasma mass spectroscopy
BF*XRF | lithium borate fusion with X-ray fluorescence spectroscopy
FA*AAS | fire assay with atomic absorption spectroscopy
FA*GRAV | fire assay with gravimetric finish
FA*OES | fire assay with inductively coupled plasma optical emission spectroscopy
INAA | instrumental neutron activation analysis
IRC | infrared combustion
LOI*TGA | loss on ignition by thermogravimetric analysis
PF*MS | sodium peroxide fusion with inductively coupled plasma mass spectroscopy
PF*OES | sodium peroxide fusion with inductively coupled plasma optical emission spectroscopy
PF*OES/MS | sodium peroxide fusion with ICP-OES or ICP-MS finish
Text Values:
NA | Unable to report due to QC failure (Lab 1.20)
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 627
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
AGAT Laboratories, Calgary, Alberta, Canada
Alex Stewart International, Mendoza, Argentina
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Malaga, WA, Australia
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
ARGETEST Mineral Processing, Ankara, Central Anatolia, Turkey
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Ankara, Central Anatolia, Turkey
CERTIMIN, Lima, Peru
ESAN Istanbul, Istanbul, Turkey
Inspectorate (BV), Lima, Peru
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Townsville, QLD, Australia
Intertek Testing Services Philippines, Cupang, Muntinlupa, Philippines
Koza Gold (Ovacik Gold Mine), Bergama, Izmir, Turkey
MSALABS, Vancouver, BC, Canada
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Ankara, Anatolia, Turkey
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico SA de CV, Cd. Industrial, Durango, Mexico
SGS del Peru, Lima, Peru
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Skyline Assayers & Laboratories, Tucson, Arizona, USA
Stewart Assay & Environmental Laboratories LLC, Kara-Balta, Chüy, Kyrgyzstan

## Homogeneity
ORE - Lab-Upscaled RSD Results for CRM: OREAS 627 (Execution: 1) - Analyte Au - (Gold) by INAA
Branch | LabSlot | LabCode | Round | Smp# | Lot | SID | Response SmplMass(g) | Response (ppm) | Upscaled Value (ppm)
2 | 1 | ANSLu | 1 | 10 | 5 | 255620 | 0.087918 | 1.9494 | 1.9588
2 | 1 | ANSLu | 1 | 17 | 9 | 255621 | 0.084952 | 1.9753 | 1.9602
2 | 1 | ANSLu | 1 | 20 | 10 | 255622 | 0.083885 | 1.9529 | 1.959
2 | 1 | ANSLu | 1 | 13 | 7 | 255623 | 0.085749 | 1.9081 | 1.9566
2 | 1 | ANSLu | 1 | 7 | 4 | 255624 | 0.083313 | 2.0149 | 1.9623
2 | 1 | ANSLu | 1 | 16 | 8 | 255625 | 0.088253 | 1.9282 | 1.9577
2 | 1 | ANSLu | 1 | 2 | 1 | 255626 | 0.086577 | 1.9843 | 1.9607
2 | 1 | ANSLu | 1 | 9 | 5 | 255627 | 0.0854 | 1.9458 | 1.9586
2 | 1 | ANSLu | 1 | 11 | 6 | 255628 | 0.086673 | 1.9433 | 1.9585
2 | 1 | ANSLu | 1 | 14 | 7 | 255629 | 0.085449 | 1.9311 | 1.9578
2 | 1 | ANSLu | 1 | 3 | 2 | 255630 | 0.084919 | 2.0587 | 1.9646
2 | 1 | ANSLu | 1 | 15 | 8 | 255631 | 0.083242 | 1.8964 | 1.956
2 | 1 | ANSLu | 1 | 12 | 6 | 255632 | 0.085576 | 1.9552 | 1.9591
2 | 1 | ANSLu | 1 | 1 | 1 | 255633 | 0.08267 | 1.9415 | 1.9584
2 | 1 | ANSLu | 1 | 18 | 9 | 255634 | 0.085963 | 1.9495 | 1.9588
2 | 1 | ANSLu | 1 | 5 | 3 | 255635 | 0.084088 | 1.9975 | 1.9614
2 | 1 | ANSLu | 1 | 6 | 3 | 255636 | 0.087992 | 1.9605 | 1.9594
2 | 1 | ANSLu | 1 | 19 | 10 | 255637 | 0.083206 | 1.9712 | 1.96
2 | 1 | ANSLu | 1 | 8 | 4 | 255638 | 0.08521 | 1.9586 | 1.9593
2 | 1 | ANSLu | 1 | 4 | 2 | 255639 | 0.082818 | 1.9644 | 1.9596
Mean | [6 empty] | 0.085193 | 1.9593 | 1.9593
Median | [7 empty] | 1.954 | 1.9591
SD | [7 empty] | 0.036136 | 0.0019256
RSD | [7 empty] | 1.844% | 0.09828%
RSD Upscale Factor | [7 empty] | 0.053289
Parameters | :
Nominal Mass Value |  | 30 | gm

## Fire Assay
| Analytical results for Ag in OREAS 627 (Indicative Value 57 ppm) | [31 empty] | Indicative
Ag | Round | Replicate | Lab | [29 empty] | 1
| No | No | 1.28 | [29 empty] | ppm
|  |  | FA*GRAV | [29 empty] | 1
|  |  | 30g | [29 empty] | 1
| 1 | 1 | 56 | [29 empty] | 1
| 1 | 2 | 57 | [29 empty] | 1
| 1 | 3 | 58 | [29 empty] | 16
| 1 | 4 | 57 | [29 empty] | 57
| 1 | 5 | 59 | [29 empty] | 7
| 1 | 6 | 55
| Mean |  | 57
| Median |  | 57
| Std Dev. |  | 1.4142
| Rel.Std.Dev. |  | 2.481%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Au in OREAS 627 (Certified Value 1.88 ppm) | [31 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 2
|  |  | 0.085g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 50g | 30g | 50g | 30g | 15g | 30g | 30g | 50g | 30g | 15g | 30g | 30g | 50g | 30g | 50g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 1.9415 | 2.01 | 1.94 | 1.911 | 1.865 | 1.88 | 1.9 | 1.865 | 1.94 | 1.76 | 1.921 | 1.923 | 1.792 | 1.8832 | 1.982 | 1.8375 | 1.68 | 1.75 | 1.85 | 1.79 | 1.9 | 1.921 | 1.6147 | 1.89 | 1.851 | 1.808 | 1.83 | 1.913 | 1.93 | 1.76 | 1
| 1 | 2 | 1.9843 | 2.01 | 2.03 | 1.886 | 1.91 | 1.895 | 1.885 | 1.89 | 1.835 | 1.8 | 1.929 | 1.9 | 1.867 | 1.8915 | 1.951 | 1.8626 | 1.71 | 1.79 | 1.839 | 1.84 | 1.78 | 1.895 | 1.5881 | 1.94 | 1.835 | 1.801 | 1.865 | 1.934 | 1.9 | 1.83 | 0
| 1 | 3 | 2.0587 | 2.02 | 1.99 | 1.911 | 1.895 | 1.885 | 1.88 | 1.885 | 1.775 | 1.78 | 1.946 | 1.878 | 1.769 | 1.8998 | 1.962 | 1.8648 | 1.5 | 1.85 | 1.82 | 1.83 | 1.78 | 1.905 | 1.6279 | 1.91 | 1.831 | 1.777 | 1.844 | 1.919 | 1.95 | 1.79 | 16
| 1 | 4 | 1.9644 | 2.01 | 1.92 | 1.906 | 1.875 | 1.92 | 1.905 | 1.85 | 1.905 | 1.81 | 1.944 | 1.901 | 1.92 | 1.904 | 1.903 | 1.8246 | 1.56 | 1.82 | 1.842 | 1.8 | 1.84 | 1.854 | 1.5879 | 1.97 | 1.835 | 1.83 | 1.844 | 1.954 | 1.94 | 1.8 | 1.8785
| 1 | 5 | 1.9975 | 2.03 | 1.96 | 1.889 | 1.89 | 1.89 | 1.95 | 1.89 | 1.805 | 1.81 | 1.925 | 1.942 | 1.741 | 1.8855 | 1.931 | 1.8418 | 1.68 | 1.77 | 1.837 | 1.79 | 1.91 | 1.876 | 1.6107 | 1.91 | 1.846 | 1.819 | 1.866 | 1.969 | 1.96 | 1.76 | 7
| 1 | 6 | 1.9605 | 2.01 | 1.95 | 1.901 | 1.92 | 1.845 | 1.925 | 1.865 | 1.825 | 1.75 | 1.941 | 1.932 | 1.912 | 1.9018 | 1.933 | 1.8386 | 1.71 | 1.79 | 1.912 | 1.83 | 1.81 | 1.903 | 1.586 | 1.96 | 1.853 | 1.854 | 1.846 | 1.916 | 1.9 | 1.83
|  | 7 | 00: 2.0149
|  | 8 | 00: 1.9586
|  | 9 | 00: 1.9458
|  | 10 | 00: 1.9494
|  | 11 | 00: 1.9433
|  | 12 | 00: 1.9552
|  | 13 | 00: 1.9081
|  | 14 | 00: 1.9311
|  | 15 | 00: 1.8964
|  | 16 | 00: 1.9282
|  | 17 | 00: 1.9753
|  | 18 | 00: 1.9495
|  | 19 | 00: 1.9712
|  | 20 | 00: 1.9529
| Mean |  | 1.9593 | 2.015 | 1.965 | 1.9007 | 1.8925 | 1.8858 | 1.9075 | 1.8742 | 1.8475 | 1.785 | 1.9343 | 1.9127 | 1.8335 | 1.8943 | 1.9437 | 1.845 | 1.64 | 1.795 | 1.85 | 1.8133 | 1.8367 | 1.8923 | 1.6026 | 1.93 | 1.8418 | 1.8148 | 1.8492 | 1.9342 | 1.93 | 1.795
| Median |  | 1.954 | 2.01 | 1.955 | 1.9035 | 1.8925 | 1.8875 | 1.9025 | 1.875 | 1.83 | 1.79 | 1.935 | 1.912 | 1.8295 | 1.8957 | 1.942 | 1.8402 | 1.68 | 1.79 | 1.8405 | 1.815 | 1.825 | 1.899 | 1.5994 | 1.925 | 1.8405 | 1.8135 | 1.845 | 1.9265 | 1.935 | 1.795
| Std Dev. |  | 0.036136 | 0.0083666 | 0.03937 | 0.010893 | 0.020676 | 0.024376 | 0.02622 | 0.016558 | 0.06259 | 0.025884 | 0.010652 | 0.02383 | 0.076427 | 0.0087971 | 0.027522 | 0.015646 | 0.088318 | 0.035637 | 0.031931 | 0.022509 | 0.057504 | 0.02383 | 0.017629 | 0.031623 | 0.009347 | 0.026271 | 0.013891 | 0.022868 | 0.025298 | 0.031464
| Rel.Std.Dev. |  | 1.844% | 0.4152% | 2.004% | 0.5731% | 1.093% | 1.293% | 1.375% | 0.8835% | 3.388% | 1.45% | 0.5507% | 1.246% | 4.168% | 0.4644% | 1.416% | 0.8481% | 5.385% | 1.985% | 1.726% | 1.241% | 3.131% | 1.259% | 1.1% | 1.638% | 0.5075% | 1.448% | 0.7512% | 1.182% | 1.311% | 1.753%
| PDM3 |  | 4.304% | 7.268% | 4.606% | 1.181% | 0.7464% | 0.3915% | 1.545% | -0.2296% | -1.649% | -4.976% | 2.973% | 1.82% | -2.394% | 0.8425% | 3.47% | -1.784% | -12.7% | -4.444% | -1.516% | -3.468% | -2.226% | 0.7375% | -14.69% | 2.743% | -1.951% | -3.388% | -1.56% | 2.964% | 2.743% | -4.444%
| Z-Score (Absolute) |  | NA | 2.34 | 1.51 | 0.44 | 0.3 | 0.19 | 0.55 | 0 | 0.44 | 1.48 | 1 | 0.64 | 0.67 | 0.33 | 1.15 | 0.48 | 3.88 | 1.31 | 0.4 | 1.01 | 0.62 | 0.3 | 4.5 | 0.93 | 0.54 | 0.98 | 0.41 | 0.99 | 0.93 | 1.31
| Analytical results for Pd in OREAS 627 (Indicative Value 0.667 ppb) | [31 empty] | Indicative
Pd | Round | Replicate | 00: Lab | 1
| No | No | 00: 1.23 | ppb
|  |  | 00: FA*OES | 3
|  |  | 00: 30g | 3
| 1 | 1 | 00: < 1 | 1
| 1 | 2 | [30 empty] | 1 | bar: 00
| 1 | 3 | 00: < 1 | 16
| 1 | 4 | 00: < 1 | 0.66667
| 1 | 5 | [30 empty] | 7 | bar: 00
| 1 | 6 | 00: < 1
Mean | bar: 00
Median | bar: 00
| Std Dev. |  | 00: 0
Rel.Std.Dev. | bar: 00
| PDM3 |  | 00: 50%
Z-Score (Absolute) | bar: 00
| Analytical results for Pt in OREAS 627 (Indicative Value < 10 ppb) | [31 empty] | Indicative
Pt | Round | Replicate | 00: Lab | 1
| No | No | 00: 1.23 | ppb
|  |  | 00: FA*OES | 1
|  |  | 00: 30g | 1
| 1 | 1 | 00: < 10 | 1
| 1 | 2 | 00: < 10 | 10
| 1 | 3 | 00: < 10 | 16
| 1 | 4 | 00: < 10 | < 10
| 1 | 5 | 00: < 10 | 7
| 1 | 6 | 00: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## AR Digest 10-50g
| Analytical results for Au in OREAS 627 (Certified Value 1.82 ppm) | [22 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.14 | 1.15 | 1.16 | 1.17 | 1.19 | 1.24 | 1.26 | 1.27 | ppm
|  |  | INAA | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*AAS | AR*AAS | AR*MS | AR*MS | AR*OES/MS | 2
|  |  | 0.085g | 30g | 30g | 15g | 25g | 25g | 25g | 25g | 25g | 25g | 25g | 15g | 10g | 50g | 25g | 10g | 30g | 50g | 30g | 25g | 20g | 3
| 1 | 1 | 1.9415 | 1.91 | 0.268 | 1.9109 | 1.77 | 1.84 | 1.9 | 1.92 | 1.66 | 1.78 | 1.943 | 1.877 |  | 1.935 | 1.8488 | 1.66 | 1.7 | 1.74 | 1.783 | 1.815 | 1.8436 | 1
| 1 | 2 | 1.9843 | 1.9 | 0.308 | 2.0508 | 1.7 | 1.83 | 1.95 | 1.83 | 1.7 | 1.75 | 1.91 | 1.8414 | 1.91 | 1.962 | 1.8886 | 1.67 | 1.7 | 1.76 | 1.797 | 1.839 | 1.8517 | 0
| 1 | 3 | 2.0587 | 1.92 | 0.195 | 1.8777 | 1.72 | 1.91 | 1.94 | 1.96 | 1.68 | 1.75 | 1.893 | 1.7866 | 1.81 | 1.924 | 1.8559 | 1.47 | 1.7 | 1.75 | 1.81 | 1.825 | 1.7727 | 16
| 1 | 4 | 1.9644 | 1.93 | 0.312 | 1.843 | 1.81 | 1.85 | 1.88 | 1.89 | 1.68 | 1.76 | 1.878 | 1.8361 |  | 1.947 | 1.8874 | 1.51 | 1.7 | 1.75 | 1.8 | 1.812 | 1.8335 | 1.8207
| 1 | 5 | 1.9975 | 1.92 | 0.284 | 1.9012 | 1.74 | 1.83 | 1.96 | 1.91 | 1.49 | 1.72 | 1.908 | 1.897 | 2 | 1.92 | 1.8892 | 1.61 | 1.7 | 1.75 | 1.799 | 1.828 | 1.7647 | 9
| 1 | 6 | 1.9605 | 1.92 | 0.147 | 1.9303 | 1.67 | 1.87 | 1.9 | 1.85 | 1.58 | 1.72 | 1.897 | 1.7732 | 1.9 | 1.889 | 1.9003 | 1.62 | 1.7 | 1.75 | 1.819 | 1.843 | 1.7428
|  | 7 | 00: 2.0149
|  | 8 | 00: 1.9586
|  | 9 | 00: 1.9458
|  | 10 | 00: 1.9494
|  | 11 | 00: 1.9433
|  | 12 | 00: 1.9552
|  | 13 | 00: 1.9081
|  | 14 | 00: 1.9311
|  | 15 | 00: 1.8964
|  | 16 | 00: 1.9282
|  | 17 | 00: 1.9753
|  | 18 | 00: 1.9495
|  | 19 | 00: 1.9712
|  | 20 | 00: 1.9529
| Mean |  | 1.9593 | 1.9167 | 0.25233 | 1.919 | 1.735 | 1.855 | 1.9217 | 1.8933 | 1.6317 | 1.7467 | 1.9048 | 1.8352 | 1.905 | 1.9295 | 1.8784 | 1.59 | 1.7 | 1.75 | 1.8013 | 1.827 | 1.8015
| Median |  | 1.954 | 1.92 | 0.276 | 1.9061 | 1.73 | 1.845 | 1.92 | 1.9 | 1.67 | 1.75 | 1.9025 | 1.8388 | 1.905 | 1.9295 | 1.888 | 1.615 | 1.7 | 1.75 | 1.7995 | 1.8265 | 1.8031
| Std Dev. |  | 0.036136 | 0.010328 | 0.066767 | 0.071222 | 0.0501 | 0.030822 | 0.032506 | 0.04761 | 0.081097 | 0.023381 | 0.021977 | 0.048604 | 0.077675 | 0.025114 | 0.0208 | 0.081731 | 0 | 0.0063246 | 0.012242 | 0.012442 | 0.046789
| Rel.Std.Dev. |  | 1.844% | 0.5388% | 26.46% | 3.711% | 2.888% | 1.662% | 1.692% | 2.515% | 4.97% | 1.339% | 1.154% | 2.648% | 4.077% | 1.302% | 1.107% | 5.14% | 0% | 0.3614% | 0.6796% | 0.681% | 2.597%
| PDM3 |  | 7.613% | 5.27% | -86.14% | 5.397% | -4.708% | 1.883% | 5.544% | 3.988% | -10.38% | -4.067% | 4.62% | 0.7961% | 4.629% | 5.974% | 3.166% | -12.67% | -6.63% | -3.884% | -1.065% | 0.3449% | -1.056%
| Z-Score (Absolute) |  | NA | 0.7 | 12.86 | 0.72 | 0.78 | 0.19 | 0.74 | 0.51 | 1.62 | 0.69 | 0.6 | 0.03 | 0.6 | 0.8 | 0.38 | 1.96 | 1.07 | 0.66 | 0.24 | 0.03 | 0.24
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 627 (Certified Value 62.8 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 65 | 54.8 | 65.5 | 67 | 63.1 | 63.4 | 63.7 | 64.4 | 63.7 | 62.54 | 61.7 | 60.51 | 61.1237 | 67.14 | 63.0621 | 63.53 | 63.3 | 59.5 | 60.82 | 60.4 | 61.1 | 68.81 | 62.56 | 63.2 | 62.7 | > 50 | 62.61 | 63 | 61.2 | 1
| 1 | 2 | 64 | 55.4 | 66.2 | 66.5 | 65.1 | 63 | 65 | 66.5 | 66.4 | 61.96 | 63.1 | 62.69 | 60.7218 | 61.96 | 61.8638 | 61.13 | 66.1 | 59.3 | 60.58 | 60.4 | 63.4 | 67.23 | 61.98 | 62.8 | 62.8 | > 50 | 63.35 | 63 | 61.8 | 1
| 1 | 3 | 64 | 56.5 | 65.5 | 66.3 | 62.5 | 61.7 | 63 | 64.7 | 62.8 | 63.2 | 59.2 | 61.02 | 60.1249 | 61.77 | 62.5368 | 64.1 | 63.9 | 59.7 | 60.1 | 62.1 | 60.7 | 69.15 | 61.6 | 63.5 | 62.2 | > 50 | 62.02 | 62 | 62 | 16
| 1 | 4 | 66 | 55.5 | 65.8 | 65.2 | 62.8 | 63.6 | 63.2 | 65.6 | 63.1 | 61.92 | 61.6 | 63.4 | 61.6969 | 61.88 | 62.0157 | 63.15 | 64.2 | 59 | 60.74 | 61.4 | 60.3 | 67.03 | 63.07 | 63.5 | 62.7 | > 50 | 62.7 | 62 | 60.3 | 62.7682
| 1 | 5 | 65 | 55.9 | 65.9 | 65.5 | 64 | 63.1 | 62.4 | 65.3 | 65.6 | 61.29 | 60.9 | 62.72 | 62.0691 | 61.48 | 62.7043 | 63.33 | 63.3 | 59.2 | 60.77 | 60.6 | 62.4 | 67.63 | 63.18 | 62.7 | 63 | > 50 | 62.31 | 62 | 62.3 | 11
| 1 | 6 | 65 | 55.6 | 65 | 65.2 | 63.7 | 63 | 62.9 | 67.1 | 66.3 | 62.45 | 63.2 | 62.73 | 61.1436 | 59.69 | 63.1578 | 63.78 | 64.2 | 59.2 | 60.81 | 60.2 | 60.7 | 68.71 | 63.02 | 62.5 | 62.4 | > 50 | 61.92 | 63 | 61.7
| Mean |  | 64.8333 | 55.6167 | 65.65 | 65.95 | 63.5333 | 62.9667 | 63.3667 | 65.6 | 64.65 | 62.2267 | 61.6167 | 62.1783 | 61.1467 | 62.32 | 62.5568 | 63.17 | 64.1667 | 59.3167 | 60.6367 | 60.85 | 61.4333 | 68.0933 | 62.5683 | 63.0333 | 62.6333 |  | 62.485 | 62.5 | 61.55
| Median |  | 65 | 55.55 | 65.65 | 65.9 | 63.4 | 63.05 | 63.1 | 65.45 | 64.65 | 62.205 | 61.65 | 62.705 | 61.1337 | 61.825 | 62.6205 | 63.43 | 64.05 | 59.25 | 60.755 | 60.5 | 60.9 | 68.17 | 62.79 | 63 | 62.7 |  | 62.46 | 62.5 | 61.75
| Std Dev. |  | 0.75277 | 0.56362 | 0.41352 | 0.75565 | 0.94798 | 0.66533 | 0.9048 | 1.0392 | 1.638 | 0.6539 | 1.4878 | 1.1382 | 0.68975 | 2.5092 | 0.53136 | 1.054 | 1.0309 | 0.24833 | 0.27689 | 0.74229 | 1.2061 | 0.90566 | 0.65018 | 0.4274 | 0.28752 |  | 0.52462 | 0.54772 | 0.71204
| Rel.Std.Dev. |  | 1.161% | 1.013% | 0.6299% | 1.146% | 1.492% | 1.057% | 1.428% | 1.584% | 2.534% | 1.051% | 2.415% | 1.83% | 1.128% | 4.026% | 0.8494% | 1.668% | 1.607% | 0.4186% | 0.4566% | 1.22% | 1.963% | 1.33% | 1.039% | 0.678% | 0.459% |  | 0.8396% | 0.8764% | 1.157%
| PDM3 |  | 3.29% | -11.39% | 4.591% | 5.069% | 1.219% | 0.3161% | 0.9534% | 4.511% | 2.998% | -0.8628% | -1.835% | -0.9398% | -2.583% | -0.7141% | -0.3369% | 0.6401% | 2.228% | -5.499% | -3.396% | -3.056% | -2.127% | 8.484% | -0.3185% | 0.4223% | -0.2149% |  | -0.4513% | -0.4274% | -1.941%
| Z-Score (Absolute) |  | 1.54 | 4.72 | 2.1 | 2.3 | 0.66 | 0.27 | 0.55 | 2.07 | 1.42 | 0.23 | 0.64 | 0.26 | 0.96 | 0.16 | 0 | 0.41 | 1.09 | 2.21 | 1.31 | 1.16 | 0.77 | 3.76 | 0 | 0.32 | 0.05 | NA | 0.05 | 0.04 | 0.69
| Analytical results for Al in OREAS 627 (Certified Value 5.49 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 5.44 | 4.98 | 5.54 | 5.28 | 5.54 | 4.84 | 5.2 | 5.48 | 5.66 | 5.47 | 5.7644 | 4.6065 | 5.5085 | 5.27 | 5.49 | 5.9 | 5.9833 | 5.1957 | 5.32 | 5.566 | 5.37 | 5.35 | 5.49 | 5.64 | 5.89 | 5.37 |  |  |  | 1
| 1 | 2 | 5.43 | 4.96 | 5.6 | 5.49 | 5.73 | 4.64 | 5.39 | 5.75 | 5.62 | 5.6 | 5.7226 | 4.5785 | 5.3474 | 5.15 | 5.44 | 6.04 | 5.9754 | 5.2369 | 5.72 | 5.5243 | 5.37 | 5.36 | 5.51 | 5.58 | 5.86 | 5.75 |  |  |  | 0
| 1 | 3 | 5.49 | 5.13 | 5.51 | 5.37 | 5.52 | 4.79 | 5.34 | 5.54 | 5.48 | 5.33 | 5.6799 | 4.5961 | 5.4141 | 5.23 | 5.56 | 5.91 | 5.9288 | 5.2965 | 5.56 | 5.5025 | 5.37 | 5.39 | 5.5 | 5.56 | 5.81 | 5.77 |  |  |  | 16
| 1 | 4 | 5.56 | 4.98 | 5.59 | 5.35 | 5.62 | 4.77 | 5.27 | 5.5 | 5.55 | 5.45 | 5.6747 | 4.8016 | 5.512 | 5.16 | 5.54 | 6.09 | 5.9846 | 5.3107 | 5.17 | 5.5561 | 5.38 | 5.35 | 5.58 | 5.55 | 5.87 | 5.32 |  |  |  | 5.4886
| 1 | 5 | 5.44 | 5.16 | 5.49 | 5.42 | 5.37 | 4.77 | 5.15 | 5.57 | 5.67 | 5.45 | 5.7385 | 4.5594 | 5.4333 | 5.25 | 5.39 | 6 | 5.9734 | 5.2256 | 5.63 | 5.5215 | 5.4 | 5.34 | 5.6 | 5.54 | 5.9 | 5.78 |  |  |  | 12
| 1 | 6 | 5.48 | 4.91 | 5.35 | 5.31 | 5.55 | 4.6 | 5.13 | 5.58 | 5.9 | 5.46 | 5.6534 | 4.7425 | 5.3876 | 5.17 | 5.5 | 6.07 | 5.9801 | 5.2017 | 5.46 | 5.5674 | 5.44 | 5.49 | 5.58 | 5.58 | 5.89 | 5.69
| Mean |  | 5.4733 | 5.02 | 5.5133 | 5.37 | 5.555 | 4.735 | 5.2467 | 5.57 | 5.6467 | 5.46 | 5.7056 | 4.6474 | 5.4338 | 5.205 | 5.4867 | 6.0017 | 5.9709 | 5.2445 | 5.4767 | 5.5396 | 5.3883 | 5.38 | 5.5433 | 5.575 | 5.87 | 5.6133
| Median |  | 5.46 | 4.98 | 5.525 | 5.36 | 5.545 | 4.77 | 5.235 | 5.555 | 5.64 | 5.455 | 5.7012 | 4.6013 | 5.4237 | 5.2 | 5.495 | 6.02 | 5.9778 | 5.2313 | 5.51 | 5.5402 | 5.375 | 5.355 | 5.545 | 5.57 | 5.88 | 5.72
| Std Dev. |  | 0.048854 | 0.1006 | 0.090921 | 0.076158 | 0.11879 | 0.093541 | 0.10482 | 0.096333 | 0.14334 | 0.08579 | 0.042823 | 0.099615 | 0.065827 | 0.051284 | 0.06314 | 0.08085 | 0.021095 | 0.048404 | 0.20422 | 0.02713 | 0.027869 | 0.056569 | 0.048442 | 0.035637 | 0.032863 | 0.21078
| Rel.Std.Dev. |  | 0.8926% | 2.004% | 1.649% | 1.418% | 2.138% | 1.976% | 1.998% | 1.729% | 2.539% | 1.571% | 0.7505% | 2.143% | 1.211% | 0.9853% | 1.151% | 1.347% | 0.3533% | 0.9229% | 3.729% | 0.4897% | 0.5172% | 1.051% | 0.8739% | 0.6392% | 0.5599% | 3.755%
| PDM3 |  | -0.2784% | -8.538% | 0.4504% | -2.161% | 1.21% | -13.73% | -4.408% | 1.483% | 2.88% | -0.5213% | 3.953% | -15.33% | -0.9985% | -5.167% | -0.03546% | 9.348% | 8.788% | -4.447% | -0.2177% | 0.9296% | -1.827% | -1.979% | 0.997% | 1.574% | 6.949% | 2.272%
| Z-Score (Absolute) |  | 0.05 | 2.92 | 0.2 | 0.71 | 0.46 | 4.72 | 1.49 | 0.56 | 1.04 | 0.14 | 1.42 | 5.27 | 0.3 | 1.75 | 0.03 | 3.29 | 3.09 | 1.5 | 0.03 | 0.37 | 0.59 | 0.64 | 0.39 | 0.59 | 2.45 | 0.83
| Analytical results for As in OREAS 627 (Certified Value 309 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 316 | 277 | 315 | 317 | 294 | 305 | 319 | 324 | 336 | 313.8 | 295 | 306.8 | 285.3953 | 293.288 | 302.2021 | 240.4 | 319 | 98.4 | 272 | 285.7 | 307 | 310 | 367 | 312 | 320 | 319.5 | 322 | 310 |  | 1
| 1 | 2 | 309 | 280 | 318 | 323 | 302 | 309 | 323 | 325 | 334 | 321.8 | 307 | 321.8 | 283.7993 | 290.626 | 298.2308 | 240.3 | 323 | 99 | 277 | 286.6 | 323 | 293 | 375 | 314 | 308 | 322.4 | 326 | 310 |  | 21
| 1 | 3 | 313 | 287 | 314 | 313 | 294 | 308 | 314 | 317 | 327 | 326.5 | 279 | 313.8 | 273.4767 | 289.502 | 296.4899 | 236.8 | 320 | 89.3 | 275 | 291.3 | 306 | 294 | 377 | 314 | 321 | 310.4 | 321 | 310 |  | 16
| 1 | 4 | 318 | 281 | 318 | 312 | 294 | 312 | 314 | 320 | 323 | 316.6 | 294 | 320.2 | 284.5654 | 284.21 | 298.7019 | 238.9 | 319 | 93.6 | 268 | 292.7 | 307 | 318 | 366 | 311 | 325 | 323 | 323 | 310 |  | 309.018
| 1 | 5 | 314 | 286 | 315 | 316 | 301 | 307 | 312 | 316 | 381 | 308 | 290 | 317.2 | 291.3905 | 291.008 | 303.8523 | 243.8 | 314 | 98 | 269 | 286.9 | 321 | 303 | 377 | 313 | 317 | 316.8 | 327 | 320 |  | 13
| 1 | 6 | 313 | 288 | 311 | 313 | 299 | 310 | 313 | 320 | 336 | 322.1 | 298 | 318.1 | 291.132 | 285.312 | 300.912 | 256.4 | 318 | 96.3 | 277 | 284.7 | 318 | 305 | 362 | 315 | 336 | 316.6 | 319 | 320
| Mean |  | 313.8333 | 283.1667 | 315.1667 | 315.6667 | 297.3333 | 308.5 | 315.8333 | 320.3333 | 339.5 | 318.1333 | 293.8333 | 316.3167 | 284.9599 | 288.991 | 300.0648 | 242.7667 | 318.8333 | 95.7667 | 273 | 287.9833 | 313.6667 | 303.8333 | 370.6667 | 313.1667 | 321.1667 | 318.1167 | 323 | 313.3333
| Median |  | 313.5 | 283.5 | 315 | 314.5 | 296.5 | 308.5 | 314 | 320 | 335 | 319.2 | 294.5 | 317.65 | 284.9803 | 290.064 | 299.807 | 240.35 | 319 | 97.15 | 273.5 | 286.75 | 312.5 | 304 | 371 | 313.5 | 320.5 | 318.15 | 322.5 | 310
| Std Dev. |  | 3.0605 | 4.446 | 2.6394 | 4.0825 | 3.7771 | 2.429 | 4.2622 | 3.6148 | 21.0024 | 6.6782 | 9.2394 | 5.4039 | 6.5269 | 3.5174 | 2.7434 | 7.0594 | 2.9269 | 3.7195 | 3.9497 | 3.2351 | 7.8401 | 9.5376 | 6.4704 | 1.472 | 9.2394 | 4.6409 | 3.0332 | 5.164
| Rel.Std.Dev. |  | 0.9752% | 1.57% | 0.8375% | 1.293% | 1.27% | 0.7874% | 1.35% | 1.128% | 6.186% | 2.099% | 3.144% | 1.708% | 2.29% | 1.217% | 0.9143% | 2.908% | 0.918% | 3.884% | 1.447% | 1.123% | 2.499% | 3.139% | 1.746% | 0.47% | 2.877% | 1.459% | 0.9391% | 1.648%
| PDM3 |  | 1.558% | -8.366% | 1.99% | 2.152% | -3.781% | -0.1676% | 2.205% | 3.662% | 9.864% | 2.95% | -4.914% | 2.362% | -7.785% | -6.481% | -2.897% | -21.44% | 3.176% | -69.01% | -11.66% | -6.807% | 1.504% | -1.678% | 19.95% | 1.343% | 3.931% | 2.944% | 4.525% | 1.396%
| Z-Score (Absolute) |  | 0.03 | 2.38 | 0.13 | 0.17 | 1.27 | 0.39 | 0.18 | 0.54 | 2.04 | 0.36 | 1.55 | 0.22 | 2.24 | 1.93 | 1.06 | 5.56 | 0.42 | 17.11 | 3.18 | 2 | 0.01 | 0.76 | 4.49 | 0.03 | 0.6 | 0.36 | 0.75 | 0.01
| Analytical results for B in OREAS 627 (Indicative Value 36.8 ppm) | [30 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.02: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02 |  | ppm
|  |  | 1.01: 4A*OES; 1.02: 4A*MS |  | 1
1
| 1 | 1 | 1.01: 36; 1.02: < 20 |  | 1
| 1 | 2 | 1.01: 36; 1.02: < 20 |  | 3
| 1 | 3 | 1.01: 37; 1.02: < 20 |  | 16
| 1 | 4 | 1.01: 38; 1.02: < 20 |  | 36.8333
| 1 | 5 | 1.01: 37; 1.02: < 20 |  | 9
| 1 | 6 | 1.01: 37; 1.02: < 20
| Mean |  | 1.01: 36.8333
| Median |  | 1.01: 37
| Std Dev. |  | 1.01: 0.75277
| Rel.Std.Dev. |  | 1.01: 2.044%
| PDM3 |  | 1.01: 8.882e-14%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.67
| Analytical results for Ba in OREAS 627 (Indicative Value 185 ppm) | [30 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | [7 empty] | 0
0
| 1 | 1 | 107 | 52 | 91 | 510 | 1970 | 30 | 200 | 30 | 115 | 24 | 4029.8861 | 4032 | 1050 | > 1000 | 412 | 583 | 54 | 133 | 1129 | 927 | 113 | 220 | [7 empty] | 1
| 1 | 2 | 116 | 53 | 75 | 570 | 1970 | 30 | 170 | 30 | 151 | 18 | 3383.6047 | 4164 | 1160 | > 1000 | 408 | 553 | 52 | 138 | 1114 | 993 | 110 | 304 | [7 empty] | 4
| 1 | 3 | 166 | 34 | 68 | 420 | 1970 | 30 | 280 | 50 | 122 | 22 | 3863.1054 | 4153 | 1250 | > 1000 | 403 | 613 | 61 | 134 | 1147 | 1014 | 94 | 300 | [7 empty] | 16
| 1 | 4 | 213 | 66 | 88 | 430 | 2600 | 60 | 860 | 20 | 127 | 24 | 3643.2199 | 4099 | 1240 | > 1000 | 415 | 618 | 65 | 133 | 1151 | 928 | 109 | 231 | [7 empty] | 184.7044
| 1 | 5 | 197 | 50 | 61 | 660 | 3280 | 40 | 140 | 40 | 129 | 22 | 3547.7809 | 4113 | 1010 | > 1000 | 413 | 745 | 59 | 138 | 1138 | 1067 | 75 | 206 | [7 empty] | 10
| 1 | 6 | 143 | 42 | 59 | 530 | 1310 | 70 | 160 | 30 | 98 | 21 | 3783.2627 | 4084 | 1100 | > 1000 | 416 | 622 | 51 | 138 | 1137 | 1310 | 122 | 221
| Mean |  | 157 | 49.5 | 73.6667 | 520 | 2183.3333 | 43.3333 | 301.6667 | 33.3333 | 123.6667 | 21.8333 | 3708.4766 | 4107.5 | 1135 |  | 411.1667 | 622.3333 | 57 | 135.6667 | 1136 | 1039.8333 | 103.8333 | 247
| Median |  | 154.5 | 51 | 71.5 | 520 | 1970 | 35 | 185 | 30 | 124.5 | 22 | 3713.2413 | 4106 | 1130 |  | 412.5 | 615.5 | 56.5 | 136 | 1137.5 | 1003.5 | 109.5 | 226
| Std Dev. |  | 42.8812 | 10.8397 | 13.5302 | 89.8888 | 674.6159 | 17.5119 | 277.8789 | 10.328 | 17.4547 | 2.2286 | 231.7025 | 48.2027 | 98.9444 |  | 4.8751 | 65.5856 | 5.5498 | 2.582 | 13.2966 | 142.7073 | 16.774 | 43.359
| Rel.Std.Dev. |  | 27.31% | 21.9% | 18.37% | 17.29% | 30.9% | 40.41% | 92.11% | 30.98% | 14.11% | 10.21% | 6.248% | 1.174% | 8.718% |  | 1.186% | 10.54% | 9.736% | 1.903% | 1.17% | 13.72% | 16.15% | 17.55%
| PDM3 |  | -15% | -73.2% | -60.12% | 181.5% | 1082% | -76.54% | 63.32% | -81.95% | -33.05% | -88.18% | 1908% | 2124% | 514.5% |  | 122.6% | 236.9% | -69.14% | -26.55% | 515% | 463% | -43.78% | 33.73%
| Z-Score (Absolute) |  | 0.3 | 0.65 | 0.57 | 0.9 | 6.41 | 0.67 | 0.18 | 0.71 | 0.41 | 0.75 | 11.46 | 12.78 | 2.94 | NA | 0.54 | 1.24 | 0.63 | 0.37 | 2.94 | 2.62 | 0.47 | 0
| Analytical results for Be in OREAS 627 (Certified Value 1.81 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | < 5 | 1.8 | 2 | 1.82 | 1.76 | 1.69 | 1.82 | 1.77 | 2.11 | 1.81 | 2 | 1.69 | 1.594 | 1.7184 | < 2 | 1.7 | 2 | < 5 | 1.73 | 1.9 | 2.5956 | 2.08 | 2 | 2 | 2 | 1.87 | 1.9 | 1.84 |  | 1
| 1 | 2 | < 5 | 1.9 | 2 | 1.84 | 1.8 | 1.68 | 1.85 | 1.84 | 1.93 | 1.98 | 2 | 1.79 | 1.6135 | 1.7669 | < 2 | 1.8 | 2 | < 5 | 1.74 | 1.9 | 2.5003 | 2.07 | 2 | 2 | 1.9 | 1.87 | 1.7 | 1.9 |  | 14
| 1 | 3 | < 5 | 1.9 | 2 | 1.83 | 1.74 | 1.72 | 1.8 | 1.78 | 1.84 | 1.98 | 2 | 1.71 | 1.6465 | 1.6897 | < 2 | 1.8 | 2 | < 5 | 1.77 | 1.9 | 2.5593 | 1.97 | 2 | 2 | 2 | 1.84 | 1.7 | 1.85 |  | 16
| 1 | 4 | < 5 | 1.9 | 2 | 1.78 | 1.76 | 1.72 | 1.8 | 1.78 | 2.03 | 1.91 | 2 | 1.72 | 1.6135 | 1.7009 | < 2 | 1.8 | 2 | < 5 | 1.78 | 1.9 | 2.3405 | 2.09 | 2 | 2 | 2 | 1.88 | 1.7 | 1.83 |  | 1.806
| 1 | 5 | < 5 | 1.9 | 2 | 1.81 | 1.75 | 1.72 | 1.8 | 1.79 | 2 | 1.83 | 2 | 1.77 | 1.5565 | 1.5709 | < 2 | 1.8 | 2 | < 5 | 1.74 | 1.9 | 2.5185 | 2.06 | 2 | 2 | 2 | 1.84 | 1.7 | 2.08 |  | 14
| 1 | 6 | < 5 | 1.7 | 2 | 1.79 | 1.78 | 1.69 | 1.77 | 1.79 | 1.88 | 1.8 | 2 | 1.79 | 1.6165 | 1.5996 | < 2 | 1.8 | 2 | < 5 | 1.75 | 1.9 | 2.2892 | 1.97 | 2 | 2 | 2.1 | 1.85 | 1.8 | 1.88
| Mean |  |  | 1.85 | 2 | 1.8117 | 1.765 | 1.7033 | 1.8067 | 1.7917 | 1.965 | 1.885 | 2 | 1.745 | 1.6068 | 1.6744 |  | 1.7833 | 2 |  | 1.7517 | 1.9 | 2.4672 | 2.04 | 2 | 2 | 2 | 1.8583 | 1.75 | 1.8967
| Median |  |  | 1.9 | 2 | 1.815 | 1.76 | 1.705 | 1.8 | 1.785 | 1.965 | 1.87 | 2 | 1.745 | 1.6135 | 1.6953 |  | 1.8 | 2 |  | 1.745 | 1.9 | 2.5094 | 2.065 | 2 | 2 | 2 | 1.86 | 1.7 | 1.865
| Std Dev. |  |  | 0.083666 | 0 | 0.023166 | 0.021679 | 0.018619 | 0.026583 | 0.024833 | 0.10055 | 0.083126 | 0 | 0.043704 | 0.029829 | 0.07447 |  | 0.040825 | 0 |  | 0.019408 | 2.4324e-16 | 0.12363 | 0.055136 | 0 | 0 | 0.063246 | 0.017224 | 0.083666 | 0.093524
| Rel.Std.Dev. |  |  | 4.522% | 0% | 1.279% | 1.228% | 1.093% | 1.471% | 1.386% | 5.117% | 4.41% | 0% | 2.505% | 1.856% | 4.448% |  | 2.289% | 0% |  | 1.108% | 1.28e-14% | 5.011% | 2.703% | 0% | 0% | 3.162% | 0.9269% | 4.781% | 4.931%
| PDM3 |  |  | 2.436% | 10.74% | 0.3131% | -2.271% | -5.685% | 0.03625% | -0.7943% | 8.803% | 4.374% | 10.74% | -3.378% | -11.03% | -7.287% |  | -1.256% | 10.74% |  | -3.009% | 5.204% | 36.61% | 12.96% | 10.74% | 10.74% | 10.74% | 2.897% | -3.101% | 5.02%
| Z-Score (Absolute) |  | 5.46 | 0.3 | NA | 0 | 0.37 | 0.86 | 0.04 | 0.16 | 1.22 | 0.58 | NA | 0.53 | 1.63 | 1.09 | 6.44 | 0.22 | NA | 5.46 | 0.48 | 0.7 | 5.2 | 1.81 | NA | NA | 1.49 | 0.37 | 0.49 | 0.67
Results from laboratories 1.03, 1.11, 1.18, 1.24 and 1.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Bi in OREAS 627 (Certified Value 44.6 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 1
2
| 1 | 1 | 56 | 43.2 | 43.4 | 46.9 | 45.6 | 43.5 | 45.1 | 48.4 | 48.3 | 48.35 | 45.2 | 43.6 | 40.1286 | 37.377 | 40.3212 | 20.46 | 47.4 | 42.4 | 31.1 | 43.2 | 43.57 | 44.47 | 51 | 43.3 | 39.1 | 46.33 | 45 | 45 |  | 1
| 1 | 2 | 51 | 43.2 | 42.8 | 47.3 | 47.3 | 43.8 | 46.2 | 49.1 | 48.4 | 47.65 | 45 | 45.65 | 37.1071 | 35.354 | 40.1415 | 15.58 | 47.7 | 42.2 | 31.05 | 43.09 | 45.55 | 43.78 | 51 | 43.9 | 38.5 | 47.72 | 48 | 42 |  | 35
| 1 | 3 | 54 | 42.8 | 40.4 | 47.6 | 47.5 | 41.6 | 43.7 | 47.8 | 46.6 | 49.16 | 42.9 | 44.85 | 39.5893 | 35.352 | 40.4648 | 12.3 | 48.3 | 42.1 | 29.79 | 43.09 | 42.84 | 42.13 | 50 | 45.6 | 39.6 | 46.7 | 45 | 47 |  | 16
| 1 | 4 | 51 | 44 | 42.7 | 46.3 | 45.6 | 44 | 43.9 | 46.7 | 47.2 | 45.59 | 44.6 | 46.3 | 38.8505 | 33.278 | 41.2898 | 20.14 | 46.7 | 41.6 | 31.11 | 44.6 | 43.75 | 43 | 50 | 43.2 | 40.6 | 47.35 | 46 | 43 |  | 44.5835
| 1 | 5 | 51 | 42.7 | 41.8 | 46.8 | 43.7 | 40.4 | 43.6 | 46.4 | 48.8 | 46.49 | 43.4 | 44.81 | 39.4771 | 35.674 | 41.0115 | 22.08 | 47.3 | 42.3 | 31.15 | 43.87 | 44.88 | 44.93 | 51 | 44.2 | 39.6 | 47.24 | 46 | 45 |  | 15
| 1 | 6 | 54 | 43.2 | 44 | 46.2 | 46.6 | 43.7 | 43.1 | 47.5 | 47.7 | 44.91 | 46 | 46.45 | 38.5355 | 33.569 | 40.7394 | 12.6 | 47.2 | 40.1 | 31.08 | 43.22 | 44.17 | 42.88 | 50 | 45.2 | 40.9 | 45.36 | 47 | 42
| Mean |  | 52.8333 | 43.1833 | 42.5167 | 46.85 | 46.05 | 42.8333 | 44.2667 | 47.65 | 47.8333 | 47.025 | 44.5167 | 45.2767 | 38.948 | 35.1007 | 40.6614 | 17.1933 | 47.4333 | 41.7833 | 30.88 | 43.5117 | 44.1267 | 43.5317 | 50.5 | 44.2333 | 39.7167 | 46.7833 | 46.1667 | 44
| Median |  | 52.5 | 43.2 | 42.75 | 46.85 | 46.1 | 43.6 | 43.8 | 47.65 | 48 | 47.07 | 44.8 | 45.25 | 39.1638 | 35.353 | 40.6021 | 17.86 | 47.35 | 42.15 | 31.09 | 43.21 | 43.96 | 43.39 | 50.5 | 44.05 | 39.6 | 46.97 | 46 | 44
| Std Dev. |  | 2.137 | 0.45789 | 1.2719 | 0.54681 | 1.4068 | 1.4787 | 1.157 | 1.0173 | 0.82624 | 1.6446 | 1.1635 | 1.0746 | 1.0634 | 1.5037 | 0.43546 | 4.2633 | 0.53541 | 0.87044 | 0.53501 | 0.60839 | 0.9698 | 1.0554 | 0.54772 | 0.98522 | 0.90203 | 0.85385 | 1.169 | 2
| Rel.Std.Dev. |  | 4.045% | 1.06% | 2.991% | 1.167% | 3.055% | 3.452% | 2.614% | 2.135% | 1.727% | 3.497% | 2.614% | 2.374% | 2.73% | 4.284% | 1.071% | 24.8% | 1.129% | 2.083% | 1.733% | 1.398% | 2.198% | 2.424% | 1.085% | 2.227% | 2.271% | 1.825% | 2.532% | 4.545%
| PDM3 |  | 18.5% | -3.14% | -4.636% | 5.084% | 3.289% | -3.926% | -0.7106% | 6.878% | 7.289% | 5.476% | -0.1498% | 1.555% | -12.64% | -21.27% | -8.797% | -61.44% | 6.392% | -6.281% | -30.74% | -2.404% | -1.025% | -2.359% | 13.27% | -0.7854% | -10.92% | 4.934% | 3.551% | -1.309%
| Z-Score (Absolute) |  | 2.33 | 0.27 | 0.45 | 0.72 | 0.5 | 0.36 | 0.02 | 0.94 | 0.99 | 0.77 | 0.09 | 0.3 | 1.41 | 2.45 | 0.95 | 7.28 | 0.88 | 0.65 | 3.59 | 0.18 | 0.01 | 0.17 | 1.7 | 0.01 | 1.2 | 0.7 | 0.54 | 0.05
| Analytical results for Ca in OREAS 627 (Certified Value 0.776 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.74 | 0.7 | 0.5 | 0.81 | 0.78 | 0.76 | 0.74 | 0.79 | 0.75 | 0.7755 | 0.77 | 0.7606 | 0.71964 | 0.419 | 0.76733 | 0.77 | 0.829 | 0.86 | 0.8705 | 0.7585 | 0.75 | 0.77604 | 0.786 | 0.87 | 0.75 | 0.75 | 0.79 | 0.826 | 0.73 | 1
| 1 | 2 | 0.74 | 0.71 | 0.54 | 0.82 | 0.8 | 0.75 | 0.77 | 0.78 | 0.75 | 0.7811 | 0.77 | 0.8088 | 0.70237 | 0.375 | 0.754 | 0.75 | 0.848 | 0.84 | 0.8846 | 0.7604 | 0.79 | 0.7756 | 0.773 | 0.87 | 0.76 | 0.72 | 0.77 | 0.803 | 0.79 | 0
| 1 | 3 | 0.76 | 0.75 | 0.57 | 0.82 | 0.77 | 0.75 | 0.75 | 0.76 | 0.72 | 0.7667 | 0.73 | 0.8049 | 0.70394 | 0.269 | 0.75017 | 0.8 | 0.843 | 0.85 | 0.8859 | 0.7759 | 0.77 | 0.77482 | 0.768 | 0.82 | 0.79 | 0.74 | 0.79 | 0.798 | 0.81 | 16
| 1 | 4 | 0.75 | 0.74 | 0.61 | 0.81 | 0.79 | 0.75 | 0.75 | 0.76 | 0.74 | 0.7814 | 0.74 | 0.8189 | 0.71762 | 0.387 | 0.7773 | 0.74 | 0.832 | 0.85 | 0.8832 | 0.7947 | 0.73 | 0.77279 | 0.773 | 0.82 | 0.77 | 0.74 | 0.81 | 0.806 | 0.75 | 0.77627
| 1 | 5 | 0.73 | 0.76 | 0.62 | 0.81 | 0.75 | 0.74 | 0.75 | 0.76 | 0.75 | 0.7867 | 0.75 | 0.8058 | 0.71757 | 0.306 | 0.77055 | 0.79 | 0.841 | 0.85 | 0.8816 | 0.7819 | 0.77 | 0.77808 | 0.773 | 0.88 | 0.78 | 0.73 | 0.79 | 0.814 | 0.81 | 16
| 1 | 6 | 0.73 | 0.7 | 0.55 | 0.79 | 0.8 | 0.73 | 0.74 | 0.78 | 0.76 | 0.7837 | 0.75 | 0.812 | 0.73265 | 0.329 | 0.77531 | 0.73 | 0.851 | 0.88 | 0.8821 | 0.7827 | 0.75 | 0.78341 | 0.796 | 0.81 | 0.78 | 0.74 | 0.8 | 0.82 | 0.8
| Mean |  | 0.74167 | 0.72667 | 0.565 | 0.81 | 0.78167 | 0.74667 | 0.75 | 0.77167 | 0.745 | 0.77918 | 0.75167 | 0.80183 | 0.71563 | 0.3475 | 0.76578 | 0.76333 | 0.84067 | 0.855 | 0.88132 | 0.77568 | 0.76 | 0.77679 | 0.77817 | 0.845 | 0.77167 | 0.73667 | 0.79167 | 0.81117 | 0.78167
| Median |  | 0.74 | 0.725 | 0.56 | 0.81 | 0.785 | 0.75 | 0.75 | 0.77 | 0.75 | 0.78125 | 0.75 | 0.8073 | 0.71759 | 0.352 | 0.76894 | 0.76 | 0.842 | 0.85 | 0.88265 | 0.7789 | 0.76 | 0.77582 | 0.773 | 0.845 | 0.775 | 0.74 | 0.79 | 0.81 | 0.795
| Std Dev. |  | 0.01169 | 0.026583 | 0.045056 | 0.010954 | 0.019408 | 0.010328 | 0.010954 | 0.013292 | 0.013784 | 0.0071399 | 0.016021 | 0.020825 | 0.011188 | 0.055956 | 0.011235 | 0.028048 | 0.0086872 | 0.013784 | 0.0055322 | 0.013989 | 0.020976 | 0.0036701 | 0.01061 | 0.031464 | 0.01472 | 0.010328 | 0.013292 | 0.010704 | 0.033714
| Rel.Std.Dev. |  | 1.576% | 3.658% | 7.974% | 1.352% | 2.483% | 1.383% | 1.461% | 1.722% | 1.85% | 0.9163% | 2.131% | 2.597% | 1.563% | 16.1% | 1.467% | 3.674% | 1.033% | 1.612% | 0.6277% | 1.803% | 2.76% | 0.4725% | 1.363% | 3.724% | 1.908% | 1.402% | 1.679% | 1.32% | 4.313%
| PDM3 |  | -4.458% | -6.391% | -27.22% | 4.344% | 0.6946% | -3.814% | -3.385% | -0.5936% | -4.029% | 0.3747% | -3.17% | 3.292% | -7.812% | -55.23% | -1.352% | -1.667% | 8.295% | 10.14% | 13.53% | -0.07619% | -2.097% | 0.06637% | 0.2437% | 8.853% | -0.5936% | -5.102% | 1.983% | 4.495% | 0.6946%
| Z-Score (Absolute) |  | 0.81 | 1.21 | 5.57 | 1.03 | 0.27 | 0.67 | 0.58 | 0 | 0.72 | 0.2 | 0.54 | 0.81 | 1.51 | 11.44 | 0.16 | 0.22 | 1.86 | 2.25 | 2.96 | 0.11 | 0.31 | 0.14 | 0.18 | 1.98 | 0 | 0.94 | 0.54 | 1.07 | 0.27
| Analytical results for Cd in OREAS 627 (Certified Value 259 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 288 | 254 | 234.8 | 271 | 248 | 249 | 259 | 266 | 280 | 269.39 | 265.2 | 249.68 | 248.225 | 272.196 | 262.3237 | 256.9 | 252 | 244.2 | 224 | 253.72 | 259 | 268.04 | 270 | 294.4 | 542 | 243.1 | 248.52 | 274 | 260 | 1
| 1 | 2 | 285 | 255 | 239.3 | 273 | 254 | 245 | 264 | 273 | 275 | 264.9 | 269.6 | 258.23 | 246.6763 | 270.576 | 256.5848 | 255.6 | 259 | 249 | 220 | 252.05 | 272 | 264.04 | 265 | 296.8 | 574 | 247.2 | 259.33 | 276 | 260 | 0
| 1 | 3 | 291 | 256 | 220.1 | 273 | 247 | 245 | 254 | 267 | 266 | 273.95 | 256.3 | 251.97 | 247.6058 | 272.374 | 260.7695 | 258.3 | 257 | 246.2 | 221 | 259.39 | 258 | 269.62 | 254 | 295.1 | 213 | 244.2 | 242.55 | 270 | 260 | 16
| 1 | 4 | 293 | 257 | 232.5 | 265 | 246 | 251 | 252 | 266 | 267 | 264.57 | 263.8 | 263.81 | 252.3102 | 262.76 | 259.8756 | 253.3 | 257 | 242.2 | 223 | 258.22 | 259 | 261.5 | 264 | 288.3 | 387 | 246.7 | 248.49 | 271 | 260 | 259.3345
| 1 | 5 | 292 | 257 | 227.5 | 269 | 254 | 248 | 255 | 265 | 281 | 258.5 | 266.4 | 258.24 | 249.7747 | 268.067 | 257.1264 | 257.6 | 256 | 245.5 | 222 | 249.95 | 270 | 264.76 | 272 | 295.3 | 563 | 243.2 | 247.33 | 273 | 270 | 17
| 1 | 6 | 288 | 257 | 231.3 | 264 | 250 | 245 | 253 | 265 | 283 | 268.7 | 270.5 | 259.87 | 248.5596 | 265.059 | 259.954 | 255.5 | 257 | 245.8 | 225 | 253.14 | 269 | 270.11 | 255 | 290.3 | 208 | 252.5 | 244.55 | 274 | 260
| Mean |  | 289.5 | 256 | 230.9167 | 269.1667 | 249.8333 | 247.1667 | 256.1667 | 267 | 275.3333 | 266.6683 | 265.3 | 256.9667 | 248.8586 | 268.5053 | 259.439 | 256.2 | 256.3333 | 245.4833 | 222.5 | 254.4117 | 264.5 | 266.345 | 263.3333 | 293.3667 | 414.5 | 246.15 | 248.4617 | 273 | 261.6667
| Median |  | 289.5 | 256.5 | 231.9 | 270 | 249 | 246.5 | 254.5 | 266 | 277.5 | 266.8 | 265.8 | 258.235 | 248.3923 | 269.3215 | 259.9148 | 256.25 | 257 | 245.65 | 222.5 | 253.43 | 264 | 266.4 | 264.5 | 294.75 | 464.5 | 245.45 | 247.91 | 273.5 | 260
| Std Dev. |  | 3.0166 | 1.2649 | 6.5831 | 3.92 | 3.4881 | 2.5626 | 4.535 | 3.0332 | 7.3394 | 5.269 | 5.099 | 5.2257 | 1.9787 | 3.9485 | 2.1928 | 1.7956 | 2.3381 | 2.2543 | 1.8708 | 3.657 | 6.473 | 3.4372 | 7.4744 | 3.3068 | 171.9008 | 3.5647 | 5.825 | 2.1909 | 4.0825
| Rel.Std.Dev. |  | 1.042% | 0.4941% | 2.851% | 1.456% | 1.396% | 1.037% | 1.77% | 1.136% | 2.666% | 1.976% | 1.922% | 2.034% | 0.7951% | 1.471% | 0.8452% | 0.7008% | 0.9121% | 0.9183% | 0.8408% | 1.437% | 2.447% | 1.291% | 2.838% | 1.127% | 41.47% | 1.448% | 2.344% | 0.8025% | 1.56%
| PDM3 |  | 11.63% | -1.286% | -10.96% | 3.791% | -3.664% | -4.692% | -1.222% | 2.956% | 6.169% | 2.828% | 2.3% | -0.913% | -4.04% | 3.536% | 0.04031% | -1.209% | -1.157% | -5.341% | -14.2% | -1.898% | 1.992% | 2.703% | 1.542% | 13.12% | 59.83% | -5.084% | -4.193% | 5.269% | 0.8993%
| Z-Score (Absolute) |  | 2.24 | 0.26 | 2.12 | 0.72 | 0.71 | 0.91 | 0.24 | 0.56 | 1.18 | 0.54 | 0.44 | 0.18 | 0.79 | 0.67 | 0 | 0.24 | 0.23 | 1.04 | 2.75 | 0.37 | 0.38 | 0.51 | 0.29 | 2.52 | 11.53 | 0.99 | 0.82 | 1.01 | 0.17
| Analytical results for Ce in OREAS 627 (Certified Value 53 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 0
0
| 1 | 1 | 36.3 | 48 | 54.1 | 57 | 37.9 | 47.8 | 51 | 65.9 | 69.78 | 25 | 63 | 43.6303 | 66 | 44.56 | 48.7 | 48.9 | 49.51 | 50.1 | 44.71 | 63.78 | 53 | [8 empty] | 1
| 1 | 2 | 37.4 | 49 | 51.6 | 59.5 | 36.2 | 49.8 | 52.4 | 69.7 | 70.51 | 26 | 64 | 42.329 | 68 | 45.53 | 49.5 | 54.3 | 52.45 | 53.1 | 46.07 | 62.11 | 64 | [8 empty] | 36
| 1 | 3 | 37.8 | 47 | 52 | 55.3 | 36.6 | 51.2 | 53.3 | 74.9 | 71.65 | 24 | 60 | 43.2085 | 68 | 45.84 | 50.1 | 52.8 | 51.04 | 62.3 | 44.82 | 64.85 | 62 | [8 empty] | 16
| 1 | 4 | 36.1 | 48 | 54.5 | 61.1 | 38.1 | 47 | 49.3 | 64.6 | 68.34 | 29 | 59 | 45.9946 | 67 | 45.68 | 50.1 | 56.6 | 54.6 | 34 | 47.51 | 62.95 | 58 | [8 empty] | 53.2716
| 1 | 5 | 32.9 | 46 | 56.8 | 62.1 | 36.5 | 44.1 | 48.9 | 76 | 69.8 | 28 | 59 | 41.9415 | 66 | 45.77 | 49.1 | 54.3 | 57.46 | 32.9 | 46.52 | 66.27 | 64 | [8 empty] | 18
| 1 | 6 | 37 | 45 | 56.2 | 56.2 | 36.5 | 48 | 51.1 | 69.3 | 68.13 | 24 | 60 | 43.6524 | 63 | 45.36 | 48.1 | 53 | 51.98 | 61.4 | 46.41 | 65.33 | 63
| Mean |  | 36.25 | 47.1667 | 54.2 | 58.5333 | 36.9667 | 47.9833 | 51 | 70.0667 | 69.7017 | 26 | 60.8333 | 43.4594 | 66.3333 | 45.4567 | 49.2667 | 53.3167 | 52.84 | 48.9667 | 46.0067 | 64.215 | 60.6667
| Median |  | 36.65 | 47.5 | 54.3 | 58.25 | 36.55 | 47.9 | 51.05 | 69.5 | 69.79 | 25.5 | 60 | 43.4194 | 66.5 | 45.605 | 49.3 | 53.65 | 52.215 | 51.6 | 46.24 | 64.315 | 62.5
| Std Dev. |  | 1.7627 | 1.472 | 2.1194 | 2.7746 | 0.81404 | 2.4384 | 1.7065 | 4.6159 | 1.3259 | 2.0976 | 2.137 | 1.424 | 1.8619 | 0.47196 | 0.79415 | 2.553 | 2.8165 | 12.9064 | 1.0751 | 1.555 | 4.3665
| Rel.Std.Dev. |  | 4.863% | 3.121% | 3.91% | 4.74% | 2.202% | 5.082% | 3.346% | 6.588% | 1.902% | 8.068% | 3.513% | 3.277% | 2.807% | 1.038% | 1.612% | 4.788% | 5.33% | 26.36% | 2.337% | 2.422% | 7.198%
| PDM3 |  | -31.95% | -11.46% | 1.743% | 9.877% | -30.61% | -9.927% | -4.264% | 31.53% | 30.84% | -51.19% | 14.19% | -18.42% | 24.52% | -14.67% | -7.518% | 0.08453% | -0.8103% | -8.081% | -13.64% | 20.54% | 13.88%
| Z-Score (Absolute) |  | 1.32 | 0.34 | 0.29 | 0.67 | 1.26 | 0.27 | 0 | 1.71 | 1.67 | 2.24 | 0.88 | 0.67 | 1.37 | 0.5 | 0.16 | 0.21 | 0.16 | 0.18 | 0.45 | 1.18 | 0.87
| Analytical results for Co in OREAS 627 (Certified Value 7.54 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 8 | 7.2 | 7.1 | 7.8 | 8.3 | 7 | 7.5 | 7.4 | 8.24 | 7.4 | 7.7 | 8 | 6.5629 | 6.459 | 7.3756 | 13.12 | 8 | 7 | 8.53 | 8.1 | 7.62 | 9.86 | 7.2 | 11 | 7 | 7.3 | 8.4 | 7 | 8 | 1
| 1 | 2 | 8 | 6.9 | 7.3 | 7.7 | 8.5 | 7.1 | 7.6 | 7.2 | 8.4 | 7.5 | 7.9 | 8 | 6.6963 | 6.944 | 7.3164 | 13.53 | 7 | 7.3 | 8.65 | 7.9 | 7.9 | 9.14 | 7.4 | 11 | 7.1 | 7 | 8.3 | 7 | 8 | 37
| 1 | 3 | 8 | 7 | 6.9 | 7.8 | 8.4 | 6.7 | 7.4 | 7.2 | 7.98 | 7.8 | 7.4 | 8 | 6.5596 | 6.647 | 7.3698 | 13.81 | 8 | 7.3 | 8.81 | 8.1 | 7.61 | 10.29 | 7.1 | 11 | 7.2 | 7.5 | 8.4 | 7 | 8 | 16
| 1 | 4 | 8 | 7.3 | 7.1 | 7.8 | 8.2 | 7.1 | 7.3 | 7.3 | 8.13 | 7.6 | 7.7 | 8 | 6.7125 | 6.381 | 7.2771 | 13.35 | 8 | 7.1 | 8.66 | 8.1 | 7.65 | 9.35 | 7.6 | 11 | 6.9 | 7.6 | 8.4 | 7 | 7 | 7.5366
| 1 | 5 | 8 | 7 | 7 | 7.5 | 7.1 | 7 | 6.9 | 7.2 | 8.58 | 7.4 | 7.8 | 8 | 6.7165 | 6.712 | 7.2197 | 13.51 | 8 | 7.3 | 8.72 | 7.8 | 7.97 | 10.5 | 7.4 | 11 | 7.1 | 7.4 | 8.1 | 7 | 8 | 19
| 1 | 6 | 8 | 7.2 | 6.9 | 7.8 | 8.2 | 7.5 | 7.3 | 7.3 | 8.47 | 7.5 | 7.9 | 8 | 6.6642 | 6.394 | 7.2504 | 13.07 | 7 | 7.4 | 8.75 | 7.6 | 7.9 | 10.04 | 7.3 | 11 | 7.1 | 7.9 | 8.2 | 7 | 8
| Mean |  | 8 | 7.1 | 7.05 | 7.7333 | 8.1167 | 7.0667 | 7.3333 | 7.2667 | 8.3 | 7.5333 | 7.7333 | 8 | 6.652 | 6.5895 | 7.3015 | 13.3983 | 7.6667 | 7.2333 | 8.6867 | 7.9333 | 7.775 | 9.8633 | 7.3333 | 11 | 7.0667 | 7.45 | 8.3 | 7 | 7.8333
| Median |  | 8 | 7.1 | 7.05 | 7.8 | 8.25 | 7.05 | 7.35 | 7.25 | 8.32 | 7.5 | 7.75 | 8 | 6.6803 | 6.553 | 7.2967 | 13.43 | 8 | 7.3 | 8.69 | 8 | 7.775 | 9.95 | 7.35 | 11 | 7.1 | 7.45 | 8.35 | 7 | 8
| Std Dev. |  | 0 | 0.15492 | 0.15166 | 0.12111 | 0.51153 | 0.2582 | 0.24221 | 0.08165 | 0.22459 | 0.15055 | 0.18619 | 0 | 0.072672 | 0.22032 | 0.063685 | 0.27817 | 0.5164 | 0.15055 | 0.096885 | 0.20656 | 0.16502 | 0.53016 | 0.17512 | 0 | 0.10328 | 0.30166 | 0.12649 | 0 | 0.40825
| Rel.Std.Dev. |  | 0% | 2.182% | 2.151% | 1.566% | 6.302% | 3.654% | 3.303% | 1.124% | 2.706% | 1.999% | 2.408% | 0% | 1.092% | 3.344% | 0.8722% | 2.076% | 6.736% | 2.081% | 1.115% | 2.604% | 2.122% | 5.375% | 2.388% | 0% | 1.462% | 4.049% | 1.524% | 0% | 5.212%
| PDM3 |  | 6.149% | -5.793% | -6.456% | 2.611% | 7.697% | -6.235% | -2.697% | -3.581% | 10.13% | -0.04303% | 2.611% | 6.149% | -11.74% | -12.57% | -3.119% | 77.78% | 1.726% | -4.024% | 15.26% | 5.264% | 3.164% | 30.87% | -2.697% | 45.95% | -6.235% | -1.149% | 10.13% | -7.12% | 3.938%
| Z-Score (Absolute) |  | NA | 0.62 | 0.7 | 0.38 | 0.99 | 0.67 | 0.25 | 0.36 | 1.28 | 0.07 | 0.38 | 0.81 | 1.33 | 1.43 | 0.3 | 9.37 | NA | 0.41 | 1.9 | 0.7 | 0.45 | 3.76 | 0.25 | NA | 0.67 | 0.07 | 1.28 | NA | NA
Results from laboratories 1.01, 1.17, 1.24 and 1.28 were removed due to their 1 ppm reading resolution. Results from laboratory 1.30 were removed due to their 10 ppm reading resolution.
| Analytical results for Cr in OREAS 627 (Certified Value 18.1 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 21 | 18 | 14 | 19 | 18 | 20 | 17 | 18 | 15.4 | 21 | 13 | 20 | 17.0352 | 23.821 | 21.6574 | 21.09 | 20 | 20 | 17.03 | 18.4 | 18 | 15 | 14 | 16 | 19 | 16 | 17 |  |  | 1
| 1 | 2 | 21 | 16 | 13 | 18 | 17 | 19 | 16 | 19 | 16.5 | 22 | 16 | 20 | 18.5641 | 22.129 | 24.3004 | 20.6 | 20 | 21 | 17.21 | 18.2 | 19 | 16 | 14 | 16 | 17 | 15 | 17 |  |  | 15
| 1 | 3 | 21 | 18 | 13 | 19 | 16 | 19 | 16 | 19 | 14.6 | 22 | 13 | 20 | 17.7875 | 20.855 | 23.9883 | 23.46 | 22 | 24 | 17.34 | 18.1 | 18 | 16 | 14 | 16 | 16 | 17 | 18 |  |  | 16
| 1 | 4 | 22 | 19 | 11 | 18 | 17 | 22 | 16 | 19 | 15.6 | 21 | 15 | 20 | 18.1105 | 41.302 | 21.0614 | 21.01 | 20 | 21 | 17.35 | 17.8 | 19 | 19 | 15 | 17 | 16 | 18 | 14 |  |  | 18.1099
| 1 | 5 | 22 | 19 | 11 | 18 | 18 | 20 | 16 | 19 | 15.6 | 21 | 14 | 20 | 18.2355 | 19.432 | 25.1001 | 20.01 | 21 | 20 | 17.32 | 17.7 | 20 | 18 | 15 | 17 | 17 | 16 | 18 |  |  | 20
| 1 | 6 | 23 | 19 | 10 | 19 | 15 | 20 | 16 | 19 | 17.1 | 21 | 14 | 20 | 17.347 | 19.345 | 20.9225 | 20.43 | 20 | 18 | 17.38 | 17.6 | 20 | 19 | 15 | 17 | 17 | 15 | 14
| Mean |  | 21.6667 | 18.1667 | 12 | 18.5 | 16.8333 | 20 | 16.1667 | 18.8333 | 15.8 | 21.3333 | 14.1667 | 20 | 17.8466 | 24.4807 | 22.8384 | 21.1 | 20.5 | 20.6667 | 17.2717 | 17.9667 | 19 | 17.1667 | 14.5 | 16.5 | 17 | 16.1667 | 16.3333
| Median |  | 21.5 | 18.5 | 12 | 18.5 | 17 | 20 | 16 | 19 | 15.6 | 21 | 14 | 20 | 17.949 | 21.492 | 22.8229 | 20.805 | 20 | 20.5 | 17.33 | 17.95 | 19 | 17 | 14.5 | 16.5 | 17 | 16 | 17
| Std Dev. |  | 0.8165 | 1.169 | 1.5492 | 0.54772 | 1.169 | 1.0954 | 0.40825 | 0.40825 | 0.87864 | 0.5164 | 1.169 | 0 | 0.57395 | 8.4135 | 1.8329 | 1.2219 | 0.83666 | 1.9664 | 0.13197 | 0.31411 | 0.89443 | 1.7224 | 0.54772 | 0.54772 | 1.0954 | 1.169 | 1.8619
| Rel.Std.Dev. |  | 3.768% | 6.435% | 12.91% | 2.961% | 6.945% | 5.477% | 2.525% | 2.168% | 5.561% | 2.421% | 8.252% | 0% | 3.216% | 34.37% | 8.026% | 5.791% | 4.081% | 9.515% | 0.7641% | 1.748% | 4.708% | 10.03% | 3.777% | 3.32% | 6.444% | 7.231% | 11.4%
| PDM3 |  | 19.64% | 0.3134% | -33.74% | 2.154% | -7.049% | 10.44% | -10.73% | 3.995% | -12.75% | 17.8% | -21.77% | 10.44% | -1.454% | 35.18% | 26.11% | 16.51% | 13.2% | 14.12% | -4.629% | -0.791% | 4.915% | -5.208% | -19.93% | -8.89% | -6.129% | -10.73% | -9.81%
| Z-Score (Absolute) |  | 1.39 | 0.07 | 2.24 | 0.2 | 0.42 | 0.76 | 0.67 | 0.32 | 0.81 | 1.26 | 1.42 | 0.76 | 0.04 | 2.44 | 1.83 | 1.17 | 0.95 | 1.01 | 0.26 | 0 | 0.39 | 0.3 | 1.3 | 0.55 | 0.36 | 0.67 | 0.61
| Analytical results for Cs in OREAS 627 (Certified Value 4.93 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [8 empty] | 2
3
| 1 | 1 | 4.93 | 5.04 | 5.13 | 4.32 | 4.71 | 4.92 | 4.96 | 5.06 | 4.91 | 4.4223 | 4.9359 | 5.1 | 5.3 | 5.2 | 4.78 | 5.8484 | 4.95 | 3 | 4.47 | 4.95 | 5.67 | [8 empty] | 1
| 1 | 2 | 4.86 | 4.99 | 5.33 | 4.36 | 4.84 | 4.96 | 5.04 | 4.96 | 5.09 | 4.3726 | 4.8116 | 5.1 | 5.7 | 5.1 | 4.96 | 6.0071 | 4.77 | 4 | 4.46 | 4.92 | 5.97 | [8 empty] | 39
| 1 | 3 | 5.01 | 4.97 | 5.19 | 4.19 | 4.73 | 4.95 | 4.87 | 5.08 | 5.06 | 4.3667 | 4.9358 | 5.2 | 5.5 | 5.1 | 4.75 | 5.7793 | 4.78 | 4 | 4.51 | 4.9 | 5.59 | [8 empty] | 16
| 1 | 4 | 4.98 | 5.07 | 5.13 | 4.51 | 4.76 | 4.94 | 4.89 | 4.81 | 5.22 | 4.5115 | 4.906 | 5 | 5.2 | 5.3 | 4.75 | 6.0953 | 5.08 | 3 | 4.64 | 4.9 | 5.48 | [8 empty] | 4.9344
| 1 | 5 | 4.77 | 4.99 | 4.9 | 4.29 | 4.6 | 4.79 | 5.27 | 4.78 | 5.07 | 4.5097 | 5.0176 | 5.1 | 5.3 | 5.3 | 5.01 | 5.8614 | 5.02 | 3 | 4.58 | 4.92 | 5.72 | [8 empty] | 21
| 1 | 6 | 4.97 | 4.91 | 5.12 | 4.47 | 4.68 | 4.99 | 5.04 | 4.75 | 5.09 | 4.4594 | 4.8072 | 5.1 | 5.2 | 5.1 | 4.93 | 5.8574 | 4.82 | 4 | 4.67 | 4.96 | 5.51
| Mean |  | 4.92 | 4.995 | 5.1333 | 4.3567 | 4.72 | 4.925 | 5.0117 | 4.9067 | 5.0733 | 4.4404 | 4.9024 | 5.1 | 5.3667 | 5.1833 | 4.8633 | 5.9081 | 4.9033 | 3.5 | 4.555 | 4.925 | 5.6567
| Median |  | 4.95 | 4.99 | 5.13 | 4.34 | 4.72 | 4.945 | 5 | 4.885 | 5.08 | 4.4409 | 4.9209 | 5.1 | 5.3 | 5.15 | 4.855 | 5.8594 | 4.885 | 3.5 | 4.545 | 4.92 | 5.63
| Std Dev. |  | 0.089889 | 0.055767 | 0.13895 | 0.11827 | 0.08025 | 0.070071 | 0.14552 | 0.14583 | 0.098928 | 0.064127 | 0.081024 | 0.063246 | 0.19664 | 0.098319 | 0.11656 | 0.1181 | 0.13186 | 0.54772 | 0.088713 | 0.0251 | 0.17862
| Rel.Std.Dev. |  | 1.827% | 1.116% | 2.707% | 2.715% | 1.7% | 1.423% | 2.904% | 2.972% | 1.95% | 1.444% | 1.653% | 1.24% | 3.664% | 1.897% | 2.397% | 1.999% | 2.689% | 15.65% | 1.948% | 0.5096% | 3.158%
| PDM3 |  | -0.2917% | 1.228% | 4.032% | -11.71% | -4.345% | -0.1904% | 1.566% | -0.5619% | 2.816% | -10.01% | -0.6492% | 3.356% | 8.76% | 5.045% | -1.44% | 19.73% | -0.6295% | -29.07% | -7.689% | -0.1904% | 14.64%
| Z-Score (Absolute) |  | 0.02 | 0.29 | 0.87 | 2.37 | 0.86 | 0 | 0.36 | 0.08 | 0.62 | 2.02 | 0.09 | 0.73 | 1.84 | 1.08 | 0.26 | 4.1 | 0.09 | NA | 1.54 | 0 | 3.05
Results from laboratory 1.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 627 (Certified Value 0.477 wt.%) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.5047 | 0.467 | 0.48483 | 0.503 | 0.485 | 0.465 | 0.467 | 0.507 | 0.487 | 0.47038 | 0.47038 | 0.4615 | 0.4466 | 0.46689 | 0.47392 | 0.5319 | 0.474 | 0.46592 | 0.4543 | 0.49067 | 0.4877 | 0.53449 | 0.4698 | 0.495 | 0.4504 | 0.4852 | 0.47458 | 0.4539 | 0.474 | 1
| 1 | 2 | 0.4994 | 0.469 | 0.49559 | 0.511 | 0.491 | 0.465 | 0.486 | 0.519 | 0.49 | 0.4681 | 0.48715 | 0.456 | 0.44546 | 0.46934 | 0.46974 | 0.5261 | 0.474 | 0.46167 | 0.4476 | 0.4917 | 0.4869 | 0.53516 | 0.4669 | 0.4906 | 0.4505 | 0.4897 | 0.48578 | 0.4581 | 0.471 | 25
| 1 | 3 | 0.4941 | 0.471 | 0.48008 | 0.51 | 0.485 | 0.47 | 0.455 | 0.517 | 0.473 | 0.47027 | 0.46134 | 0.4627 | 0.44987 | 0.46954 | 0.47141 | 0.5245 | 0.474 | 0.45839 | 0.4415 | 0.5006 | 0.4783 | 0.53488 | 0.4726 | 0.4916 | 0.4545 | 0.488 | 0.46004 | 0.4485 | 0.474 | 16
| 1 | 4 | 0.5114 | 0.478 | 0.4868 | 0.498 | 0.481 | 0.476 | 0.458 | 0.514 | 0.478 | 0.47682 | 0.47456 | 0.4544 | 0.45212 | 0.45822 | 0.4735 | 0.521 | 0.467 | 0.46284 | 0.4481 | 0.49623 | 0.4906 | 0.53746 | 0.4683 | 0.4981 | 0.451 | 0.4878 | 0.47538 | 0.4528 | 0.467 | 0.47701
| 1 | 5 | 0.5038 | 0.47 | 0.49382 | 0.498 | 0.475 | 0.466 | 0.458 | 0.517 | 0.49 | 0.45967 | 0.46858 | 0.4529 | 0.45944 | 0.4602 | 0.46855 | 0.5238 | 0.465 | 0.45503 | 0.4516 | 0.48813 | 0.4874 | 0.53271 | 0.4779 | 0.4985 | 0.4515 | 0.484 | 0.46866 | 0.458 | 0.477 | 22
| 1 | 6 | 0.504 | 0.472 | 0.48402 | 0.492 | 0.479 | 0.469 | 0.453 | 0.504 | 0.509 | 0.47428 | 0.48048 | 0.4459 | 0.4529 | 0.45654 | 0.47633 | 0.5317 | 0.474 | 0.46205 | 0.4527 | 0.48779 | 0.4885 | 0.53928 | 0.4768 | 0.4857 | 0.4536 | 0.4908 | 0.46986 | 0.4588 | 0.472
| Mean |  | 0.5029 | 0.47117 | 0.48752 | 0.502 | 0.48267 | 0.4685 | 0.46283 | 0.513 | 0.48783 | 0.46992 | 0.47375 | 0.45557 | 0.45107 | 0.46346 | 0.47224 | 0.5265 | 0.47133 | 0.46098 | 0.4493 | 0.49252 | 0.48657 | 0.53566 | 0.47205 | 0.49325 | 0.45192 | 0.48758 | 0.47238 | 0.45502 | 0.4725
| Median |  | 0.5039 | 0.4705 | 0.48581 | 0.5005 | 0.483 | 0.4675 | 0.458 | 0.5155 | 0.4885 | 0.47032 | 0.47247 | 0.4552 | 0.45099 | 0.46355 | 0.47246 | 0.5253 | 0.474 | 0.46186 | 0.44985 | 0.49118 | 0.48755 | 0.53502 | 0.4712 | 0.4933 | 0.45125 | 0.4879 | 0.47222 | 0.45595 | 0.473
| Std Dev. |  | 0.0057827 | 0.0037639 | 0.0060029 | 0.0074565 | 0.0055737 | 0.0042308 | 0.012319 | 0.0060992 | 0.012449 | 0.0059216 | 0.009139 | 0.006135 | 0.0050476 | 0.0058201 | 0.0028862 | 0.0044249 | 0.0041793 | 0.003784 | 0.0046264 | 0.0049958 | 0.0042528 | 0.0023378 | 0.0045329 | 0.0049172 | 0.0017221 | 0.0025911 | 0.0085541 | 0.0040336 | 0.0033912
| Rel.Std.Dev. |  | 1.15% | 0.7988% | 1.231% | 1.485% | 1.155% | 0.9031% | 2.662% | 1.189% | 2.552% | 1.26% | 1.929% | 1.347% | 1.119% | 1.256% | 0.6112% | 0.8404% | 0.8867% | 0.8209% | 1.03% | 1.014% | 0.8741% | 0.4364% | 0.9603% | 0.9969% | 0.3811% | 0.5314% | 1.811% | 0.8865% | 0.7177%
| PDM3 |  | 5.427% | -1.225% | 2.204% | 5.239% | 1.186% | -1.784% | -2.972% | 7.545% | 2.269% | -1.487% | -0.6841% | -4.496% | -5.439% | -2.842% | -1% | 10.37% | -1.19% | -3.36% | -5.809% | 3.251% | 2.003% | 12.3% | -1.04% | 3.404% | -5.261% | 2.216% | -0.9703% | -4.611% | -0.9458%
| Z-Score (Absolute) |  | 1.36 | 0.05 | 0.67 | 1.32 | 0.46 | 0.17 | 0.43 | 1.81 | 0.69 | 0.11 | 0.06 | 0.75 | 0.95 | 0.4 | 0.01 | 2.41 | 0.05 | 0.51 | 1.03 | 0.9 | 0.63 | 2.82 | 0.01 | 0.93 | 0.91 | 0.68 | 0 | 0.77 | 0.01
| Analytical results for Dy in OREAS 627 (Certified Value 3.44 ppm) | [30 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | 2
3
| 1 | 1 | 1.01: 3.5; 1.02: 3.17; 1.03: 3.37; 1.04: 2.4978; 1.05: 3.4584; 1.06: 3.42 | 1
| 1 | 2 | 1.01: 3.7; 1.02: 3.21; 1.03: 3.53; 1.04: 2.3594; 1.05: 3.2682; 1.06: 3.84 | 7
| 1 | 3 | 1.01: 3.8; 1.02: 3.3; 1.03: 3.42; 1.04: 2.407; 1.05: 3.5898; 1.06: 3.5 | 16
| 1 | 4 | 1.01: 3.5; 1.02: 3.06; 1.03: 3.62; 1.04: 2.5367; 1.05: 3.7421; 1.06: 3.32 | 3.4357
| 1 | 5 | 1.01: 3.6; 1.02: 3.05; 1.03: 3.45; 1.04: 2.3815; 1.05: 3.3302; 1.06: 3.53 | 23
| 1 | 6 | 1.01: 3.7; 1.02: 3.09; 1.03: 3.48; 1.04: 2.4857; 1.05: 3.5397; 1.06: 3.39
| Mean |  | 1.01: 3.6333; 1.02: 3.1467; 1.03: 3.4783; 1.04: 2.4447; 1.05: 3.488; 1.06: 3.5
| Median |  | 1.01: 3.65; 1.02: 3.13; 1.03: 3.465; 1.04: 2.4463; 1.05: 3.499; 1.06: 3.46
| Std Dev. |  | 1.01: 0.12111; 1.02: 0.098116; 1.03: 0.087958; 1.04: 0.071637; 1.05: 0.17419; 1.06: 0.18298
| Rel.Std.Dev. |  | 1.01: 3.333%; 1.02: 3.118%; 1.03: 2.529%; 1.04: 2.93%; 1.05: 4.994%; 1.06: 5.228%
| PDM3 |  | 1.01: 5.753%; 1.02: -8.412%; 1.03: 1.242%; 1.04: -28.84%; 1.05: 1.524%; 1.06: 1.872%
| Z-Score (Absolute) |  | 1.01: 1.21; 1.02: 2.72; 1.03: 0.04; 1.04: 8.39; 1.05: 0.04; 1.06: 0.14
| Analytical results for Er in OREAS 627 (Certified Value 1.46 ppm) | [30 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | 2
2
| 1 | 1 | 1.01: 1.6; 1.02: 1.43; 1.03: 1.42; 1.04: 1.1493; 1.05: 1.6346; 1.06: 1.48 | 1
| 1 | 2 | 1.01: 1.6; 1.02: 1.48; 1.03: 1.49; 1.04: 1.0933; 1.05: 1.5915; 1.06: 1.54 | 8
| 1 | 3 | 1.01: 1.6; 1.02: 1.47; 1.03: 1.5; 1.04: 1.1541; 1.05: 1.6755; 1.06: 1.49 | 16
| 1 | 4 | 1.01: 1.5; 1.02: 1.39; 1.03: 1.51; 1.04: 1.1562; 1.05: 1.7409; 1.06: 1.42 | 1.4647
| 1 | 5 | 1.01: 1.6; 1.02: 1.36; 1.03: 1.49; 1.04: 1.1198; 1.05: 1.5783; 1.06: 1.51 | 24
| 1 | 6 | 1.01: 1.7; 1.02: 1.37; 1.03: 1.47; 1.04: 1.1664; 1.05: 1.7649; 1.06: 1.43
| Mean |  | 1.01: 1.6; 1.02: 1.4167; 1.03: 1.48; 1.04: 1.1398; 1.05: 1.6643; 1.06: 1.4783
| Median |  | 1.01: 1.6; 1.02: 1.41; 1.03: 1.49; 1.04: 1.1517; 1.05: 1.6551; 1.06: 1.485
| Std Dev. |  | 1.01: 0.063246; 1.02: 0.051251; 1.03: 0.032249; 1.04: 0.02768; 1.05: 0.077075; 1.06: 0.046224
| Rel.Std.Dev. |  | 1.01: 3.953%; 1.02: 3.618%; 1.03: 2.179%; 1.04: 2.428%; 1.05: 4.631%; 1.06: 3.127%
| PDM3 |  | 1.01: 9.235%; 1.02: -3.282%; 1.03: 1.042%; 1.04: -22.18%; 1.05: 13.62%; 1.06: 0.9282%
| Z-Score (Absolute) |  | 1.01: 0.89; 1.02: 0.46; 1.03: 0.01; 1.04: 2.5; 1.05: 1.36; 1.06: 0.01
| Analytical results for Eu in OREAS 627 (Certified Value 1.32 ppm) | [30 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS | 2
3
| 1 | 1 | 1.01: 1.3; 1.02: 1.31; 1.03: 1.33; 1.04: 0.86883; 1.05: 1.3 | 1
| 1 | 2 | 1.01: 1.33; 1.02: 1.32; 1.03: 1.41; 1.04: 0.83842; 1.05: 1.37 | 9
| 1 | 3 | 1.01: 1.3; 1.02: 1.39; 1.03: 1.37; 1.04: 0.85533; 1.05: 1.31 | 16
| 1 | 4 | 1.01: 1.31; 1.02: 1.23; 1.03: 1.41; 1.04: 0.89347; 1.05: 1.27 | 1.3242
| 1 | 5 | 1.01: 1.3; 1.02: 1.29; 1.03: 1.37; 1.04: 0.82781; 1.05: 1.33 | 25
| 1 | 6 | 1.01: 1.27; 1.02: 1.32; 1.03: 1.38; 1.04: 0.86888; 1.05: 1.26
| Mean |  | 1.01: 1.3017; 1.02: 1.31; 1.03: 1.3783; 1.04: 0.85879; 1.05: 1.3067
| Median |  | 1.01: 1.3; 1.02: 1.315; 1.03: 1.375; 1.04: 0.86208; 1.05: 1.305
| Std Dev. |  | 1.01: 0.019408; 1.02: 0.051769; 1.03: 0.029944; 1.04: 0.023629; 1.05: 0.040332
| Rel.Std.Dev. |  | 1.01: 1.491%; 1.02: 3.952%; 1.03: 2.173%; 1.04: 2.751%; 1.05: 3.087%
| PDM3 |  | 1.01: -1.699%; 1.02: -1.07%; 1.03: 4.091%; 1.04: -35.14%; 1.05: -1.322%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.45; 1.03: 9.67; 1.04: 60.4; 1.05: 0
| Analytical results for Fe in OREAS 627 (Certified Value 6.86 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 6.74 | 6.53 | 6.46 | 7.05 | 6.94 | 7.07 | 6.92 | 6.58 | 7.65 | 7.04 | 6.75 | 7.04 | 6.6522 | 7.355 | 6.9588 | 6.88 | 7.03 | 6.48 | 6.7923 | 7.0278 | 6.67 | 6.7144 | 6.86 | 7.43 | 6.8 | 6.47 | 7.03 | 7.14 | 6.89 | 1
| 1 | 2 | 6.75 | 6.53 | 6.5 | 7.09 | 7.11 | 7.01 | 7.06 | 6.85 | 7.67 | 7.01 | 6.85 | 6.86 | 6.6704 | 7.64 | 6.7291 | 6.71 | 7.08 | 6.61 | 7.1548 | 6.9815 | 6.77 | 6.7518 | 6.83 | 7.38 | 6.74 | 6.64 | 7.07 | 7.2 | 7.18 | 26
| 1 | 3 | 6.83 | 6.57 | 6.39 | 7.08 | 6.93 | 6.99 | 6.8 | 6.66 | 7.47 | 6.96 | 6.44 | 7.09 | 6.6053 | 7.267 | 6.8338 | 6.8 | 7.1 | 6.5 | 7.2474 | 7.1395 | 6.76 | 6.7716 | 6.87 | 7.3 | 6.79 | 6.39 | 6.91 | 7.08 | 6.66 | 16
| 1 | 4 | 6.87 | 6.73 | 6.47 | 6.95 | 6.93 | 7.06 | 6.76 | 6.69 | 7.53 | 6.96 | 6.64 | 6.92 | 6.7415 | 7.559 | 6.9102 | 6.73 | 7.03 | 6.68 | 7.1863 | 7.1273 | 6.51 | 6.7436 | 6.86 | 7.37 | 6.73 | 6.47 | 7.06 | 7.1 | 6.58 | 6.8599
| 1 | 5 | 6.77 | 6.55 | 6.45 | 6.99 | 6.62 | 7.03 | 6.75 | 6.71 | 7.71 | 6.94 | 6.69 | 7.05 | 6.7758 | 7.298 | 6.7786 | 6.88 | 7.02 | 6.68 | 7.1548 | 7.0577 | 6.9 | 6.7922 | 6.95 | 7.47 | 6.76 | 6.41 | 6.99 | 7.16 | 7.03 | 26
| 1 | 6 | 6.79 | 6.77 | 6.41 | 6.84 | 6.97 | 6.99 | 6.76 | 6.67 | 7.26 | 6.91 | 6.84 | 6.91 | 6.7133 | 7.4 | 6.8756 | 6.76 | 6.98 | 6.65 | 7.1968 | 7.0115 | 6.72 | 6.7022 | 6.93 | 7.49 | 6.77 | 6.52 | 6.95 | 7.24 | 6.71
| Mean |  | 6.7917 | 6.6133 | 6.4467 | 7 | 6.9167 | 7.025 | 6.8417 | 6.6933 | 7.5483 | 6.97 | 6.7017 | 6.9783 | 6.6931 | 7.4198 | 6.8477 | 6.7933 | 7.04 | 6.6 | 7.1221 | 7.0575 | 6.7217 | 6.746 | 6.8833 | 7.4067 | 6.765 | 6.4833 | 7.0017 | 7.1533 | 6.8417
| Median |  | 6.78 | 6.56 | 6.455 | 7.02 | 6.935 | 7.02 | 6.78 | 6.68 | 7.59 | 6.96 | 6.72 | 6.98 | 6.6919 | 7.3775 | 6.8547 | 6.78 | 7.03 | 6.63 | 7.1706 | 7.0427 | 6.74 | 6.7477 | 6.865 | 7.405 | 6.765 | 6.47 | 7.01 | 7.15 | 6.8
| Std Dev. |  | 0.049967 | 0.10764 | 0.040332 | 0.095079 | 0.16071 | 0.034496 | 0.12432 | 0.088694 | 0.16786 | 0.047329 | 0.15224 | 0.093256 | 0.062428 | 0.14875 | 0.084929 | 0.073666 | 0.043359 | 0.089219 | 0.16511 | 0.063828 | 0.12891 | 0.0339 | 0.046332 | 0.070616 | 0.027386 | 0.089815 | 0.063377 | 0.060222 | 0.23284
| Rel.Std.Dev. |  | 0.7357% | 1.628% | 0.6256% | 1.358% | 2.323% | 0.4911% | 1.817% | 1.325% | 2.224% | 0.679% | 2.272% | 1.336% | 0.9327% | 2.005% | 1.24% | 1.084% | 0.6159% | 1.352% | 2.318% | 0.9044% | 1.918% | 0.5025% | 0.6731% | 0.9534% | 0.4048% | 1.385% | 0.9052% | 0.8419% | 3.403%
| PDM3 |  | -0.9941% | -3.594% | -6.023% | 2.043% | 0.8281% | 2.407% | -0.2653% | -2.428% | 10.04% | 1.606% | -2.306% | 1.727% | -2.431% | 8.163% | -0.1774% | -0.9698% | 2.626% | -3.788% | 3.822% | 2.882% | -2.015% | -1.66% | 0.3421% | 7.971% | -1.383% | -5.489% | 2.067% | 4.278% | -0.2653%
| Z-Score (Absolute) |  | 0.25 | 1.03 | 1.76 | 0.67 | 0.3 | 0.78 | 0.03 | 0.68 | 3.07 | 0.54 | 0.64 | 0.57 | 0.68 | 2.51 | 0 | 0.24 | 0.84 | 1.08 | 1.2 | 0.92 | 0.55 | 0.45 | 0.16 | 2.45 | 0.36 | 1.6 | 0.67 | 1.34 | 0.03
| Analytical results for Ga in OREAS 627 (Certified Value 18.9 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 1
2
| 1 | 1 | 19 | 18.6 | 19.6 | 18.05 | 16.7 | 18 | 20 | 20.3 | 19.05 | 17.78 | 20.1002 | 19.095 | 17.3978 | 18.6 | 19 | 15.3 | 10.35 | 17.93 | 17.91 | 18.2 | 17.6 | 20.18 | 21.14 | [6 empty] | 1
| 1 | 2 | 17 | 18.6 | 19.05 | 18.9 | 17.95 | 18.65 | 20.1 | 20.4 | 19.19 | 18.33 | 20.1748 | 18.387 | 17.5212 | 18.5 | 20 | 15.17 | 12.1 | 19.16 | 18.37 | 19.4 | 17.6 | 20.72 | 22.61 | [6 empty] | 44
| 1 | 3 | 19 | 19.4 | 19.65 | 18.95 | 16.35 | 18.2 | 19.8 | 19.7 | 19.34 | 18.64 | 19.9678 | 18.163 | 17.296 | 18.4 | 19 | 14.61 | 10.62 | 18.42 | 17.12 | 19.2 | 17.8 | 20.52 | 21.53 | [6 empty] | 16
| 1 | 4 | 17 | 18.2 | 19.3 | 18.25 | 16.55 | 18.1 | 20.2 | 19.9 | 19.18 | 18.86 | 20.1383 | 17.788 | 17.0128 | 18.5 | 18 | 15.21 | 14.54 | 18.08 | 18.16 | 19.4 | 18.3 | 20.14 | 20.95 | [6 empty] | 18.8694
| 1 | 5 | 19 | 19.3 | 19.5 | 17.35 | 16.1 | 17.45 | 19.5 | 21.4 | 19.36 | 18.56 | 20.3886 | 18.332 | 17.4409 | 18.9 | 19 | 15.22 | 10.5 | 19.03 | 18.36 | 19.7 | 17.9 | 19.65 | 19.42 | [6 empty] | 27
| 1 | 6 | 19 | 19.2 | 19.05 | 18.25 | 17.55 | 18.25 | 20.1 | 20.4 | 19.32 | 18.84 | 20.4959 | 17.89 | 17 | 18.8 | 19 | 15.18 | 10.46 | 18.85 | 17.95 | 19.6 | 18.6 | 19.46 | 19.42
| Mean |  | 18.3333 | 18.8833 | 19.3583 | 18.2917 | 16.8667 | 18.1083 | 19.95 | 20.35 | 19.24 | 18.5017 | 20.2109 | 18.2758 | 17.2781 | 18.6167 | 19 | 15.115 | 11.4283 | 18.5783 | 17.9783 | 19.25 | 17.9667 | 20.1117 | 20.845
| Median |  | 19 | 18.9 | 19.4 | 18.25 | 16.625 | 18.15 | 20.05 | 20.35 | 19.255 | 18.6 | 20.1566 | 18.2475 | 17.3469 | 18.55 | 19 | 15.195 | 10.56 | 18.635 | 18.055 | 19.4 | 17.85 | 20.16 | 21.045
| Std Dev. |  | 1.0328 | 0.48339 | 0.26724 | 0.59196 | 0.72434 | 0.39168 | 0.25884 | 0.58907 | 0.12083 | 0.4039 | 0.19532 | 0.46607 | 0.2227 | 0.19408 | 0.63246 | 0.25161 | 1.6583 | 0.51184 | 0.46353 | 0.54314 | 0.40332 | 0.48582 | 1.2445
| Rel.Std.Dev. |  | 5.633% | 2.56% | 1.38% | 3.236% | 4.294% | 2.163% | 1.297% | 2.895% | 0.628% | 2.183% | 0.9664% | 2.55% | 1.289% | 1.043% | 3.329% | 1.665% | 14.51% | 2.755% | 2.578% | 2.822% | 2.245% | 2.416% | 5.97%
| PDM3 |  | -2.841% | 0.07385% | 2.591% | -3.062% | -10.61% | -4.033% | 5.727% | 7.847% | 1.964% | -1.949% | 7.11% | -3.146% | -8.433% | -1.339% | 0.6921% | -19.9% | -39.43% | -1.543% | -4.722% | 2.017% | -4.784% | 6.584% | 10.47%
| Z-Score (Absolute) |  | NA | 0.31 | 0.78 | 0.29 | 1.72 | 0.47 | 1.38 | 1.78 | 0.66 | 0.08 | 1.64 | 0.3 | 1.31 | 0.04 | NA | 3.48 | 7.18 | 0 | 0.6 | 0.67 | 0.61 | 1.54 | 2.28
Results from laboratories 1.01 and 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 627 (Certified Value 4.87 ppm) | [30 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.22; 1.07: 1.30 | [6 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: 4.5; 1.02: 4.64; 1.04: 4.75; 1.05: 3.4933; 1.06: 6.0375; 1.07: 5.19 | [6 empty] | 1
| 1 | 2 | 1.01: 4.6; 1.02: 4.7; 1.04: 5.15; 1.05: 3.3545; 1.06: 5.9793; 1.07: 5.5 | [6 empty] | 10
| 1 | 3 | 1.01: 4.5; 1.02: 4.9; 1.04: 5.03; 1.05: 3.5295; 1.06: 6.3096; 1.07: 5.21 | [6 empty] | 16
| 1 | 4 | 1.01: 4.7; 1.02: 4.68; 1.04: 5.14; 1.05: 3.7081; 1.06: 6.612; 1.07: 5.07 | [6 empty] | 4.8698
| 1 | 5 | 1.01: 4.7; 1.02: 4.53; 1.04: 5.01; 1.05: 3.4446; 1.06: 5.9926; 1.07: 5.34 | [6 empty] | 28
| 1 | 6 | 1.01: 4.6; 1.02: 4.57; 1.04: 5.15; 1.05: 3.4358; 1.06: 6.0961; 1.07: 5.16
| Mean |  | 1.01: 4.6; 1.02: 4.67; 1.04: 5.0383; 1.05: 3.4943; 1.06: 6.1712; 1.07: 5.245
| Median |  | 1.01: 4.6; 1.02: 4.66; 1.04: 5.085; 1.05: 3.469; 1.06: 6.0668; 1.07: 5.2
| Std Dev. |  | 1.01: 0.089443; 1.02: 0.12992; 1.04: 0.15446; 1.05: 0.12033; 1.06: 0.24725; 1.07: 0.15241
| Rel.Std.Dev. |  | 1.01: 1.944%; 1.02: 2.782%; 1.04: 3.066%; 1.05: 3.444%; 1.06: 4.006%; 1.07: 2.906%
| PDM3 |  | 1.01: -5.54%; 1.02: -4.103%; 1.04: 3.461%; 1.05: -28.25%; 1.06: 26.72%; 1.07: 7.705%
| Z-Score (Absolute) |  | 1.01: 0.53; 1.02: 0.39; 1.04: 0.39; 1.05: 2.84; 1.06: 2.75; 1.07: 0.82
| Analytical results for Ge in OREAS 627 (Indicative Value 0.22 ppm) | [30 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 2
2
| 1 | 1 | 0.3 | 0.26 | 0.18 | 0.14 | 0.09 | 0.18 | 0.19 | 0.3 | 2.3 | 1.8 | < 0.2 | 0.47586 | 2 | 4.1 | 0.41 | 1.3 | < 0.05 | 0.14 | [11 empty] | 1
| 1 | 2 | 0.3 | 0.38 | 0.2 | 0.15 | 0.1 | 0.17 | 0.15 | 0.3 | 2.5 | 1.8 | < 0.2 | 0.41271 | 2.3 | 4.2 | 0.3 | 1 | < 0.05 | 0.13 | [11 empty] | 46
| 1 | 3 | 0.2 | 0.35 | 0.19 | 0.15 | 0.09 | 0.18 | 0.2 | 0.3 | 2.4 | 1.8 | < 0.2 | 0.43617 | 2.4 | 4.1 | 0.37 | 1 | 0.05 | 0.11 | [11 empty] | 16
| 1 | 4 | 0.2 | 0.37 | 0.16 | 0.16 | 0.09 | 0.17 | 0.13 | 0.29 | 2.1 | 1.8 | < 0.2 | 0.40962 | 2.3 | 4.2 | 0.37 | 1.3 | 0.06 | 0.13 | [11 empty] | 0.21774
| 1 | 5 | 0.1 | 0.29 | 0.18 | 0.17 | 0.09 | 0.16 | 0.19 | 0.33 | 2.3 | 1.6 | < 0.2 | 0.43447 | 2.3 | 4.3 | 0.48 | 1.3 | 0.05 | 0.09 | [11 empty] | 9
| 1 | 6 | 0.2 | 0.32 | 0.15 | 0.14 | 0.09 | 0.19 | 0.22 | 0.3 | 2.2 | 1.7 | < 0.2 | 0.4052 | 2.2 | 4 | 0.41 | 1.2 | 0.07 | 0.11
| Mean |  | 0.21667 | 0.32833 | 0.17667 | 0.15167 | 0.091667 | 0.175 | 0.18 | 0.30333 | 2.3 | 1.75 |  | 0.429 | 2.25 | 4.15 | 0.39 | 1.1833 | 0.0575 | 0.11833
| Median |  | 0.2 | 0.335 | 0.18 | 0.15 | 0.09 | 0.175 | 0.19 | 0.3 | 2.3 | 1.8 |  | 0.42359 | 2.3 | 4.15 | 0.39 | 1.25 | 0.055 | 0.12
| Std Dev. |  | 0.075277 | 0.047081 | 0.018619 | 0.01169 | 0.0040825 | 0.010488 | 0.033466 | 0.013663 | 0.14142 | 0.083666 |  | 0.026398 | 0.13784 | 0.10488 | 0.059666 | 0.1472 | 0.0095743 | 0.018348
| Rel.Std.Dev. |  | 34.74% | 14.34% | 10.54% | 7.708% | 4.454% | 5.993% | 18.59% | 4.504% | 6.149% | 4.781% |  | 6.153% | 6.126% | 2.527% | 15.3% | 12.44% | 16.65% | 15.51%
| PDM3 |  | -0.4913% | 50.79% | -18.86% | -30.34% | -57.9% | -19.63% | -17.33% | 39.31% | 956.3% | 703.7% |  | 97.03% | 933.4% | 1806% | 79.12% | 443.5% | -73.59% | -45.65%
| Z-Score (Absolute) |  | 0.19 | 0.31 | 0.37 | 0.48 | 0.75 | 0.38 | 0.36 | 0.19 | 9.12 | 6.66 | 0.72 | 0.76 | 8.9 | 17.39 | 0.58 | 4.13 | 0.95 | 0.63
| Analytical results for Hf in OREAS 627 (Certified Value 4.38 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 2
3
| 1 | 1 | 4.4 | 4.5 | 4.8 | 4.6 | 4.2 | 4.1 | 4.3 | 4.8 | 4.17 | 4.3 | 4.64 | 3.8693 | 4.2646 | 4.4 | 4.3 | 4.46 | 4.14 | 4.6366 | 4.52 | 4.81 | 4.45 | 4.2 | 4.2 | [6 empty] | 1
| 1 | 2 | 4.4 | 4.4 | 4.8 | 4.9 | 4.1 | 4.2 | 4.6 | 5 | 4.24 | 4.3 | 4.34 | 3.6887 | 4.263 | 4.5 | 4.3 | 4.46 | 4.45 | 4.5967 | 4.31 | 4.91 | 4.47 | 4.4 | 4.3 | [6 empty] | 16
| 1 | 3 | 4.4 | 4.1 | 4.9 | 4.6 | 4 | 4.2 | 4.2 | 4.7 | 4.42 | 4 | 4.4 | 3.6828 | 4.3145 | 4.5 | 4.3 | 4.44 | 4.25 | 4.6984 | 4.14 | 4.95 | 4.52 | 4.4 | 4.2 | [6 empty] | 16
| 1 | 4 | 4.6 | 4.1 | 4.8 | 4.7 | 4.1 | 4.1 | 4.4 | 4.8 | 4.35 | 4 | 4.47 | 3.8157 | 4.3039 | 4.4 | 4.2 | 4.78 | 4.03 | 4.5313 | 4.25 | 4.99 | 4.5 | 4.5 | 4.2 | [6 empty] | 4.3802
| 1 | 5 | 4.4 | 4.2 | 4.8 | 4.4 | 3.9 | 4.1 | 4.3 | 5 | 4.4 | 4 | 4.31 | 3.7247 | 4.2875 | 4.4 | 4.3 | 4.63 | 4.26 | 4.6544 | 4.28 | 4.91 | 4.69 | 4.3 | 5 | [6 empty] | 29
| 1 | 6 | 4.5 | 4.5 | 4.5 | 4.8 | 4 | 4 | 4.4 | 4.9 | 4.56 | 4.2 | 4.51 | 3.9079 | 4.2631 | 4.6 | 4.3 | 4.54 | 4.24 | 4.5656 | 4.59 | 5.06 | 4.57 | 4.3 | 4.3
| Mean |  | 4.45 | 4.3 | 4.7667 | 4.6667 | 4.05 | 4.1167 | 4.3667 | 4.8667 | 4.3567 | 4.1333 | 4.445 | 3.7815 | 4.2828 | 4.4667 | 4.2833 | 4.5517 | 4.2283 | 4.6138 | 4.3483 | 4.9383 | 4.5333 | 4.35 | 4.3667
| Median |  | 4.4 | 4.3 | 4.8 | 4.65 | 4.05 | 4.1 | 4.35 | 4.85 | 4.375 | 4.1 | 4.435 | 3.7702 | 4.276 | 4.45 | 4.3 | 4.5 | 4.245 | 4.6166 | 4.295 | 4.93 | 4.51 | 4.35 | 4.25
| Std Dev. |  | 0.083666 | 0.18974 | 0.13663 | 0.17512 | 0.10488 | 0.075277 | 0.13663 | 0.12111 | 0.13837 | 0.15055 | 0.12178 | 0.096361 | 0.022723 | 0.08165 | 0.040825 | 0.13243 | 0.13992 | 0.061251 | 0.17151 | 0.084479 | 0.087331 | 0.10488 | 0.31411
| Rel.Std.Dev. |  | 1.88% | 4.412% | 2.866% | 3.753% | 2.59% | 1.829% | 3.129% | 2.488% | 3.176% | 3.642% | 2.74% | 2.548% | 0.5306% | 1.828% | 0.9531% | 2.909% | 3.309% | 1.328% | 3.944% | 1.711% | 1.926% | 2.411% | 7.193%
| PDM3 |  | 1.594% | -1.83% | 8.824% | 6.541% | -7.538% | -6.016% | -0.3081% | 11.11% | -0.5364% | -5.635% | 1.48% | -13.67% | -2.223% | 1.975% | -2.211% | 3.915% | -3.466% | 5.335% | -0.7267% | 12.74% | 3.497% | -0.6886% | -0.3081%
| Z-Score (Absolute) |  | 0.41 | 0.32 | 1.95 | 1.46 | 1.54 | 1.22 | 0 | 2.44 | 0.05 | 1.14 | 0.38 | 2.85 | 0.41 | 0.49 | 0.41 | 0.9 | 0.67 | 1.2 | 0.09 | 2.79 | 0.81 | 0.08 | 0
| Analytical results for Hg in OREAS 627 (Indicative Value 2.59 ppm) | [30 empty] | Indicative
Hg | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab | [6 empty] | 1
| No | No | 1.02: 1.13; 1.03: 1.14; 1.04: 1.20; 1.05: 1.30 | [6 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*OES/MS; 1.04: 4A*OES/MS; 1.05: 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | 1.02: 2.6339; 1.03: 2.7421; 1.04: 2.59; 1.05: < 2 | [6 empty] | 1
| 1 | 2 | 1.02: 2.5796; 1.03: 1.7613; 1.04: 2.58; 1.05: < 2 | [6 empty] | 4
| 1 | 3 | 1.02: 2.5639; 1.03: 2.8637; 1.04: 2.67; 1.05: < 2 | [6 empty] | 16
| 1 | 4 | 1.02: 2.5877; 1.03: 2.2344; 1.04: 2.68; 1.05: < 2 | [6 empty] | 2.591
| 1 | 5 | 1.02: 2.6449; 1.03: 2.3805; 1.04: 2.75; 1.05: < 2 | [6 empty] | 10
| 1 | 6 | 1.02: 2.6213; 1.03: 2.4844; 1.04: 2.49; 1.05: < 2
| Mean |  | 1.02: 2.6052; 1.03: 2.4111; 1.04: 2.6267
| Median |  | 1.02: 2.6045; 1.03: 2.4324; 1.04: 2.63
| Std Dev. |  | 1.02: 0.032646; 1.03: 0.39349; 1.04: 0.091797
| Rel.Std.Dev. |  | 1.02: 1.253%; 1.03: 16.32%; 1.04: 3.495%
| PDM3 |  | 1.02: 0.5494%; 1.03: -6.943%; 1.04: 1.378%
| Z-Score (Absolute) |  | 1.02: 0.61; 1.03: 0.61; 1.04: 0.74; 1.05: 9.43
| Analytical results for Ho in OREAS 627 (Certified Value 0.57 ppm) | [30 empty] | Certified
Ho | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.02: 1.02; 1.03: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.22; 1.07: 1.30 | [6 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.02: 0.6; 1.03: 0.54; 1.04: 0.55; 1.05: 0.43524; 1.06: 0.56305; 1.07: 0.59 | [6 empty] | 1
| 1 | 2 | 1.02: 0.6; 1.03: 0.52; 1.04: 0.56; 1.05: 0.43046; 1.06: 0.55361; 1.07: 0.64 | [6 empty] | 6
| 1 | 3 | 1.02: 0.6; 1.03: 0.56; 1.04: 0.56; 1.05: 0.42915; 1.06: 0.59506; 1.07: 0.6 | [6 empty] | 16
| 1 | 4 | 1.02: 0.6; 1.03: 0.52; 1.04: 0.58; 1.05: 0.44877; 1.06: 0.6231; 1.07: 0.58 | [6 empty] | 0.57484
| 1 | 5 | 1.02: 0.6; 1.03: 0.52; 1.04: 0.56; 1.05: 0.43585; 1.06: 0.55056; 1.07: 0.61 | [6 empty] | 30
| 1 | 6 | 1.02: 0.6; 1.03: 0.51; 1.04: 0.57; 1.05: 0.44827; 1.06: 0.6377; 1.07: 0.59
| Mean |  | 1.02: 0.6; 1.03: 0.52833; 1.04: 0.56333; 1.05: 0.43796; 1.06: 0.58718; 1.07: 0.60167
| Median |  | 1.02: 0.6; 1.03: 0.52; 1.04: 0.56; 1.05: 0.43555; 1.06: 0.57905; 1.07: 0.595
| Std Dev. |  | 1.02: 0; 1.03: 0.018348; 1.04: 0.010328; 1.05: 0.0085876; 1.06: 0.037294; 1.07: 0.02137
| Rel.Std.Dev. |  | 1.02: 0%; 1.03: 3.473%; 1.04: 1.833%; 1.05: 1.961%; 1.06: 6.351%; 1.07: 3.552%
| PDM3 |  | 1.02: 4.378%; 1.03: -8.09%; 1.04: -2.001%; 1.05: -23.81%; 1.06: 2.147%; 1.07: 4.668%
| Z-Score (Absolute) |  | 1.02: 0.65; 1.03: 1.24; 1.04: 0.31; 1.05: 3.62; 1.06: 0.31; 1.07: 0.7
| Analytical results for In in OREAS 627 (Certified Value 8.83 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 2
3
| 1 | 1 | 8 | 8.7 | 8.82 | 9.54 | 9.25 | 9.48 | 8.67 | 9.19 | 9.39 | 9.42 | 8.73 | 8.33 | 8.2364 | 8.58 | 8.59 | 8.858 | 9.1 | 8.328 | 7.67 | 7.89 | 8.722 | 8.894 | [7 empty] | 1
| 1 | 2 | 8 | 8.5 | 9.17 | 9.54 | 9.38 | 9.33 | 8.93 | 9.2 | 8.93 | 9.2 | 9 | 8.77 | 8.1953 | 8.95 | 8.6 | 8.803 | 9.6 | 8.395 | 8.02 | 7.86 | 8.573 | 9.209 | [7 empty] | 33
| 1 | 3 | 9 | 8.6 | 8.17 | 9.44 | 9.6 | 8.78 | 8.71 | 9.54 | 8.68 | 9.54 | 8.47 | 8.82 | 8.1985 | 8.95 | 8.51 | 8.776 | 9.2 | 8.378 | 8.14 | 8 | 8.67 | 8.588 | [7 empty] | 16
| 1 | 4 | 9 | 8.7 | 8.91 | 9.45 | 9.36 | 9.1 | 8.63 | 9.27 | 8.62 | 9.24 | 8.72 | 8.96 | 8.2351 | 8.73 | 8.58 | 9.027 | 9.2 | 8.577 | 7.69 | 8.1 | 8.728 | 8.792 | [7 empty] | 8.8261
| 1 | 5 | 8 | 8.8 | 8.72 | 9.47 | 8.59 | 9.02 | 8.56 | 9.19 | 9.27 | 8.99 | 8.75 | 8.82 | 8.2289 | 8.89 | 8.62 | 8.88 | 9.6 | 8.669 | 8.05 | 8.1 | 8.678 | 8.903 | [7 empty] | 31
| 1 | 6 | 9 | 8.7 | 8.86 | 9.31 | 9.31 | 9.62 | 8.7 | 9.47 | 9.04 | 9.21 | 8.98 | 8.92 | 8.2153 | 8.93 | 8.7 | 8.719 | 9.4 | 8.428 | 8.08 | 8.44 | 8.487 | 8.491
| Mean |  | 8.5 | 8.6667 | 8.775 | 9.4583 | 9.2483 | 9.2217 | 8.7 | 9.31 | 8.9883 | 9.2667 | 8.775 | 8.77 | 8.2183 | 8.8383 | 8.6 | 8.8438 | 9.35 | 8.4625 | 7.9417 | 8.065 | 8.643 | 8.8128
| Median |  | 8.5 | 8.7 | 8.84 | 9.46 | 9.335 | 9.215 | 8.685 | 9.235 | 8.985 | 9.225 | 8.74 | 8.82 | 8.2221 | 8.91 | 8.595 | 8.8305 | 9.3 | 8.4115 | 8.035 | 8.05 | 8.674 | 8.843
| Std Dev. |  | 0.54772 | 0.10328 | 0.33243 | 0.084715 | 0.34371 | 0.31231 | 0.12522 | 0.15556 | 0.30903 | 0.19138 | 0.19563 | 0.2268 | 0.018182 | 0.15132 | 0.061644 | 0.10669 | 0.21679 | 0.1317 | 0.20663 | 0.20974 | 0.094505 | 0.25543
| Rel.Std.Dev. |  | 6.444% | 1.192% | 3.788% | 0.8957% | 3.716% | 3.387% | 1.439% | 1.671% | 3.438% | 2.065% | 2.229% | 2.586% | 0.2212% | 1.712% | 0.7168% | 1.206% | 2.319% | 1.556% | 2.602% | 2.601% | 1.093% | 2.898%
| PDM3 |  | -3.695% | -1.806% | -0.5788% | 7.163% | 4.784% | 4.482% | -1.429% | 5.483% | 1.838% | 4.992% | -0.5788% | -0.6354% | -6.887% | 0.1388% | -2.562% | 0.2011% | 5.936% | -4.119% | -10.02% | -8.623% | -2.074% | -0.1501%
| Z-Score (Absolute) |  | NA | 0.34 | 0 | 2.16 | 1.5 | 1.41 | 0.24 | 1.69 | 0.67 | 1.55 | 0 | 0.02 | 1.76 | 0.2 | 0.55 | 0.22 | 1.82 | 0.99 | 2.63 | 2.24 | 0.42 | 0.12
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for K in OREAS 627 (Certified Value 2.51 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 2.57 | 2.13 | 2.46 | 2.65 | 2.53 | 2.51 | 2.5 | 2.59 | 2.32 | 2.5392 | 2.59 | 2.5948 | 2.3406 | 2.039 | 2.5226 | 2.58 | 2.65 | 2.54 | 2.5718 | 2.382 | 2.49 | 2.3475 | 2.44 | 2.76 | 2.52 | 2.44 | 2.63 | 2.44 | 2.56 | 1
| 1 | 2 | 2.6 | 2.31 | 2.5 | 2.69 | 2.58 | 2.49 | 2.55 | 2.71 | 2.25 | 2.5447 | 2.66 | 2.5706 | 2.3327 | 1.949 | 2.4767 | 2.52 | 2.65 | 2.53 | 2.6845 | 2.3914 | 2.61 | 2.305 | 2.46 | 2.64 | 2.54 | 2.45 | 2.63 | 2.43 | 2.95 | 0
| 1 | 3 | 2.57 | 2.47 | 2.46 | 2.67 | 2.51 | 2.48 | 2.48 | 2.64 | 2.31 | 2.5033 | 2.46 | 2.5452 | 2.3347 | 1.873 | 2.4688 | 2.56 | 2.63 | 2.52 | 2.7413 | 2.4256 | 2.53 | 2.3722 | 2.44 | 2.64 | 2.58 | 2.42 | 2.55 | 2.45 | 2.52 | 16
| 1 | 4 | 2.66 | 2.24 | 2.47 | 2.63 | 2.53 | 2.52 | 2.48 | 2.67 | 2.27 | 2.5057 | 2.56 | 2.532 | 2.3435 | 2.014 | 2.5218 | 2.52 | 2.64 | 2.53 | 2.6963 | 2.4254 | 2.43 | 2.2896 | 2.45 | 2.61 | 2.53 | 2.47 | 2.64 | 2.48 | 2.51 | 2.513
| 1 | 5 | 2.58 | 2.35 | 2.51 | 2.62 | 2.38 | 2.49 | 2.44 | 2.69 | 2.31 | 2.517 | 2.56 | 2.5601 | 2.3693 | 2.08 | 2.4699 | 2.57 | 2.66 | 2.53 | 2.6754 | 2.393 | 2.58 | 2.3038 | 2.47 | 2.68 | 2.53 | 2.41 | 2.6 | 2.48 | 2.62 | 32
| 1 | 6 | 2.64 | 2.35 | 2.44 | 2.58 | 2.53 | 2.48 | 2.44 | 2.64 | 2.29 | 2.5065 | 2.6 | 2.5381 | 2.4022 | 2.004 | 2.4864 | 2.54 | 2.65 | 2.54 | 2.6537 | 2.3812 | 2.5 | 2.3604 | 2.48 | 2.64 | 2.58 | 2.47 | 2.59 | 2.5 | 2.49
| Mean |  | 2.6033 | 2.3083 | 2.4733 | 2.64 | 2.51 | 2.495 | 2.4817 | 2.6567 | 2.2917 | 2.5194 | 2.5717 | 2.5568 | 2.3539 | 1.9932 | 2.491 | 2.5483 | 2.6467 | 2.5317 | 2.6705 | 2.3998 | 2.5233 | 2.3298 | 2.4567 | 2.6617 | 2.5467 | 2.4433 | 2.6067 | 2.4633 | 2.6083
| Median |  | 2.59 | 2.33 | 2.465 | 2.64 | 2.53 | 2.49 | 2.48 | 2.655 | 2.3 | 2.5118 | 2.575 | 2.5526 | 2.3421 | 2.009 | 2.4815 | 2.55 | 2.65 | 2.53 | 2.6799 | 2.3922 | 2.515 | 2.3262 | 2.455 | 2.64 | 2.535 | 2.445 | 2.615 | 2.465 | 2.54
| Std Dev. |  | 0.038297 | 0.11496 | 0.026583 | 0.038987 | 0.067823 | 0.016432 | 0.041191 | 0.04274 | 0.027142 | 0.018173 | 0.065853 | 0.023429 | 0.027102 | 0.072915 | 0.024954 | 0.025626 | 0.010328 | 0.0075277 | 0.056431 | 0.020497 | 0.065013 | 0.034509 | 0.01633 | 0.053072 | 0.026583 | 0.025033 | 0.033862 | 0.027325 | 0.1736
| Rel.Std.Dev. |  | 1.471% | 4.98% | 1.075% | 1.477% | 2.702% | 0.6586% | 1.66% | 1.609% | 1.184% | 0.7213% | 2.561% | 0.9163% | 1.151% | 3.658% | 1.002% | 1.006% | 0.3902% | 0.2973% | 2.113% | 0.8541% | 2.576% | 1.481% | 0.6647% | 1.994% | 1.044% | 1.025% | 1.299% | 1.109% | 6.656%
| PDM3 |  | 3.596% | -8.143% | -1.577% | 5.055% | -0.1181% | -0.715% | -1.246% | 5.718% | -8.806% | 0.256% | 2.336% | 1.744% | -6.332% | -20.68% | -0.873% | 1.407% | 5.32% | 0.7441% | 6.269% | -4.505% | 0.4125% | -7.291% | -2.24% | 5.917% | 1.341% | -2.771% | 3.729% | -1.975% | 3.795%
| Z-Score (Absolute) |  | 0.74 | 1.87 | 0.41 | 1.07 | 0.08 | 0.22 | 0.33 | 1.22 | 2.02 | 0 | 0.46 | 0.33 | 1.47 | 4.66 | 0.25 | 0.26 | 1.13 | 0.11 | 1.34 | 1.06 | 0.03 | 1.68 | 0.56 | 1.26 | 0.24 | 0.67 | 0.77 | 0.5 | 0.79
| Analytical results for La in OREAS 627 (Certified Value 22.6 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.16 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 20 | 12.7 | 17.6 | 18.6 | 24.4 | 14.9 | 19.5 | 17 | 29 | 32.38 | 8 | 29.93 | 16.2117 | 31.2 | 37 | 21.71 | 23.5 | 20.4 | 18.51 | 22 | 24.2 | 21 | 29.8 | 15 | 30.6 |  |  |  |  | 1
| 1 | 2 | 21 | 13.2 | 18.2 | 18 | 25.9 | 13.7 | 20.3 | 16.7 | 30.4 | 32.92 | 9.1 | 31.31 | 15.9081 | 31.25 | 38 | 22.16 | 23.7 | 22.8 | 18.85 | 23 | 25 | 21.4 | 27.7 | 13.5 | 34.7 |  |  |  |  | 34
| 1 | 3 | 19 | 10.8 | 17.6 | 18 | 22.8 | 13.6 | 21.9 | 18.3 | 32.3 | 33.12 | 7.7 | 30.75 | 16.1969 | 31.22 | 38 | 22.46 | 23.8 | 22.7 | 20.02 | 23 | 31.1 | 21.1 | 30.6 | 13.4 | 32.7 |  |  |  |  | 16
| 1 | 4 | 19 | 10.8 | 18.7 | 19.2 | 25.8 | 14.9 | 19 | 15.9 | 28.5 | 31.41 | 9.8 | 31.39 | 17.5213 | 30.66 | 38 | 22.17 | 23.8 | 19.5 | 21.63 | 22 | 24.3 | 22.2 | 30.9 | 11.9 | 31 |  |  |  |  | 22.5771
| 1 | 5 | 20 | 8.7 | 17 | 20 | 27.3 | 13.9 | 18.2 | 16.2 | 32.7 | 31.93 | 9 | 31 | 15.5826 | 31.05 | 35 | 22.19 | 23.5 | 23.5 | 21.81 | 23 | 22.3 | 22.1 | 31.1 | 12.7 | 33.1 |  |  |  |  | 33
| 1 | 6 | 19 | 11.7 | 16.9 | 19.7 | 24.2 | 13.3 | 19.6 | 16.8 | 28.8 | 32.2 | 8.1 | 32.21 | 16.5252 | 30.7 | 32 | 22.18 | 23.6 | 23 | 19.03 | 21 | 29.3 | 22 | 32.1 | 14.8 | 32
| Mean |  | 19.6667 | 11.3167 | 17.6667 | 18.9167 | 25.0667 | 14.05 | 19.75 | 16.8167 | 30.2833 | 32.3267 | 8.6167 | 31.0983 | 16.3243 | 31.0133 | 36.3333 | 22.145 | 23.65 | 21.9833 | 19.975 | 22.3333 | 26.0333 | 21.6333 | 30.3667 | 13.55 | 32.35
| Median |  | 19.5 | 11.25 | 17.6 | 18.9 | 25.1 | 13.8 | 19.55 | 16.75 | 29.7 | 32.29 | 8.55 | 31.155 | 16.2043 | 31.125 | 37.5 | 22.175 | 23.65 | 22.75 | 19.525 | 22.5 | 24.65 | 21.7 | 30.75 | 13.45 | 32.35
| Std Dev. |  | 0.8165 | 1.6117 | 0.69186 | 0.85421 | 1.5845 | 0.68629 | 1.2629 | 0.83287 | 1.8411 | 0.63213 | 0.8085 | 0.75653 | 0.66715 | 0.26756 | 2.4221 | 0.24189 | 0.13784 | 1.6241 | 1.4431 | 0.8165 | 3.3975 | 0.53166 | 1.5042 | 1.1946 | 1.4977
| Rel.Std.Dev. |  | 4.152% | 14.24% | 3.916% | 4.516% | 6.321% | 4.885% | 6.395% | 4.953% | 6.08% | 1.955% | 9.383% | 2.433% | 4.087% | 0.8627% | 6.666% | 1.092% | 0.5828% | 7.388% | 7.225% | 3.656% | 13.05% | 2.458% | 4.954% | 8.816% | 4.63%
| PDM3 |  | -12.89% | -49.88% | -21.75% | -16.21% | 11.03% | -37.77% | -12.52% | -25.51% | 34.13% | 43.18% | -61.83% | 37.74% | -27.7% | 37.37% | 60.93% | -1.914% | 4.752% | -2.63% | -11.53% | -1.08% | 15.31% | -4.18% | 34.5% | -39.98% | 43.29%
| Z-Score (Absolute) |  | 0.3 | 1.39 | 0.56 | 0.4 | 0.4 | 1.04 | 0.29 | 0.67 | 1.08 | 1.35 | 1.74 | 1.19 | 0.74 | 1.18 | 1.87 | 0.02 | 0.22 | 0 | 0.26 | 0.05 | 0.53 | 0.05 | 1.09 | 1.1 | 1.35
| Analytical results for Li in OREAS 627 (Certified Value 17.7 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 18 | 16.5 | 17.3 | 18.6 | 18.5 | 17.2 | 18.2 | 18.6 | 17.5 | 17.5 | 17.9 | 17.2 | 16.4454 | 18.888 | 17.143 | 15.42 | 18.4 | 15 | 18.06 | 17 | 17.3 | 24.1039 | 18.1 | 16 | 18 | 19.7 | 18 | 17.6 |  | 1
| 1 | 2 | 18 | 16.8 | 18.3 | 18.8 | 19.1 | 17.1 | 17.8 | 20.8 | 18 | 18.5 | 18.2 | 17.7 | 17.1 | 16.551 | 16.6872 | 16.87 | 18.2 | 16.5 | 17.85 | 17 | 17.3 | 24.0565 | 18.1 | 17 | 17 | 20.4 | 19 | 17.3 |  | 0
| 1 | 3 | 18 | 17.5 | 18 | 18.8 | 18.3 | 17.2 | 17.4 | 20 | 17.1 | 19 | 17.9 | 17.6 | 17.0881 | 17.56 | 17.03 | 16.26 | 18.7 | 16.7 | 17.71 | 17 | 16.4 | 24.8462 | 17.7 | 17 | 17 | 19.6 | 19 | 15.3 |  | 16
| 1 | 4 | 20 | 16.7 | 18.1 | 18.4 | 17.9 | 18 | 17.4 | 20.4 | 17.1 | 17.5 | 18.7 | 17.8 | 17.0652 | 16.843 | 16.3524 | 15.36 | 18.4 | 15.3 | 17.95 | 17 | 17.6 | 24.8545 | 18.7 | 16 | 18 | 19.8 | 19 | 16.3 |  | 17.708
| 1 | 5 | 19 | 16.8 | 18.2 | 18.4 | 17.7 | 17.8 | 17.4 | 21 | 18.3 | 17.7 | 19.1 | 17.5 | 17.4103 | 15.254 | 17.3709 | 17.01 | 18.5 | 16 | 17.82 | 17 | 17.4 | 24.417 | 18.6 | 17 | 18 | 19.9 | 19 | 14.1 |  | 34
| 1 | 6 | 19 | 16.8 | 17 | 19 | 17.9 | 17.8 | 17 | 19.3 | 18.5 | 17.8 | 19.3 | 18 | 16.786 | 16.952 | 16.5053 | 16.72 | 18.3 | 15.8 | 17.92 | 17 | 17.5 | 24.3073 | 18 | 17 | 18 | 20.1 | 18 | 15.2
| Mean |  | 18.6667 | 16.85 | 17.8167 | 18.6667 | 18.2333 | 17.5167 | 17.5333 | 20.0167 | 17.75 | 18 | 18.5167 | 17.6333 | 16.9825 | 17.008 | 16.8481 | 16.2733 | 18.4167 | 15.8833 | 17.885 | 17 | 17.25 | 24.4309 | 18.2 | 16.6667 | 17.6667 | 19.9167 | 18.6667 | 15.9667
| Median |  | 18.5 | 16.8 | 18.05 | 18.7 | 18.1 | 17.5 | 17.4 | 20.2 | 17.75 | 17.75 | 18.45 | 17.65 | 17.0767 | 16.8975 | 16.8586 | 16.49 | 18.4 | 15.9 | 17.885 | 17 | 17.35 | 24.3622 | 18.1 | 17 | 18 | 19.85 | 19 | 15.8
| Std Dev. |  | 0.8165 | 0.33912 | 0.53448 | 0.24221 | 0.5164 | 0.392 | 0.41312 | 0.92177 | 0.60581 | 0.61319 | 0.608 | 0.27325 | 0.32924 | 1.1962 | 0.39561 | 0.72948 | 0.17224 | 0.66156 | 0.12012 | 0 | 0.43243 | 0.35059 | 0.37947 | 0.5164 | 0.5164 | 0.29269 | 0.5164 | 1.3471
| Rel.Std.Dev. |  | 4.374% | 2.013% | 3% | 1.298% | 2.832% | 2.238% | 2.356% | 4.605% | 3.413% | 3.407% | 3.284% | 1.55% | 1.939% | 7.033% | 2.348% | 4.483% | 0.9352% | 4.165% | 0.6717% | 0% | 2.507% | 1.435% | 2.085% | 3.098% | 2.923% | 1.47% | 2.766% | 8.437%
| PDM3 |  | 5.413% | -4.846% | 0.6134% | 5.413% | 2.966% | -1.081% | -0.9866% | 13.04% | 0.2369% | 1.649% | 4.566% | -0.4219% | -4.097% | -3.953% | -4.856% | -8.102% | 4.002% | -10.3% | 0.9993% | -3.998% | -2.587% | 37.97% | 2.778% | -5.881% | -0.2337% | 12.47% | 5.413% | -9.834%
| Z-Score (Absolute) |  | NA | 0.82 | 0.06 | 0.83 | 0.44 | 0.21 | 0.2 | 2.06 | 0 | 0.23 | 0.7 | 0.11 | 0.7 | 0.67 | 0.82 | 1.34 | 0.61 | 1.7 | 0.12 | NA | 0.45 | 6.07 | 0.41 | NA | NA | 1.97 | NA | 1.62
Results from laboratories 1.01, 1.20, 1.25, 1.26 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 627 (Certified Value 0.19 ppm) | [30 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.23; 1.06: 1.25; 1.07: 1.26; 1.08: 1.30 |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*OES/MS; 1.07: 4A*OES/MS; 1.08: 4A*MS |  | 2
2
| 1 | 1 | 1.01: 0.2; 1.02: 0.2; 1.03: 0.2; 1.04: 0.15778; 1.05: 0.18; 1.06: 0.22; 1.07: 0.16; 1.08: 0.21 |  | 1
| 1 | 2 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.21; 1.04: 0.15777; 1.05: 0.18; 1.06: 0.22; 1.07: 0.16; 1.08: 0.23 |  | 35
| 1 | 3 | 1.01: 0.2; 1.02: 0.22; 1.03: 0.21; 1.04: 0.15945; 1.05: 0.17; 1.06: 0.23; 1.07: 0.16; 1.08: 0.21 |  | 16
| 1 | 4 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.2; 1.04: 0.16481; 1.05: 0.19; 1.06: 0.21; 1.07: 0.17; 1.08: 0.2 |  | 0.19132
| 1 | 5 | 1.01: 0.2; 1.02: 0.19; 1.03: 0.2; 1.04: 0.15519; 1.05: 0.18; 1.06: 0.21; 1.07: 0.17; 1.08: 0.22 |  | 35
| 1 | 6 | 1.01: 0.2; 1.02: 0.18; 1.03: 0.22; 1.04: 0.16029; 1.05: 0.19; 1.06: 0.23; 1.07: 0.18; 1.08: 0.21
| Mean |  | 1.01: 0.2; 1.02: 0.19167; 1.03: 0.20667; 1.04: 0.15922; 1.05: 0.18167; 1.06: 0.22; 1.07: 0.16667; 1.08: 0.21333
| Median |  | 1.01: 0.2; 1.02: 0.185; 1.03: 0.205; 1.04: 0.15862; 1.05: 0.18; 1.06: 0.22; 1.07: 0.165; 1.08: 0.21
| Std Dev. |  | 1.01: 3.0405e-17; 1.02: 0.016021; 1.03: 0.008165; 1.04: 0.0032502; 1.05: 0.0075277; 1.06: 0.0089443; 1.07: 0.008165; 1.08: 0.010328
| Rel.Std.Dev. |  | 1.01: 1.52e-14%; 1.02: 8.359%; 1.03: 3.951%; 1.04: 2.041%; 1.05: 4.144%; 1.06: 4.066%; 1.07: 4.899%; 1.08: 4.841%
| PDM3 |  | 1.01: 4.539%; 1.02: 0.183%; 1.03: 8.023%; 1.04: -16.78%; 1.05: -5.044%; 1.06: 14.99%; 1.07: -12.88%; 1.08: 11.51%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0; 1.03: 0.47; 1.04: 1.01; 1.05: 0.31; 1.06: 0.88; 1.07: 0.78; 1.08: 0.67
Results from laboratory 1.02 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Mg in OREAS 627 (Certified Value 0.382 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.39 | 0.35 | 0.37 | 0.36 | 0.37 | 0.32 | 0.36 | 0.34 | 0.4 | 0.4078 | 0.37 | 0.4072 | 0.32626 | 0.292 | 0.39533 | 0.41 | 0.394 | 0.4 | 0.3845 | 0.3991 | 0.36 | 0.41793 | 0.378 | 0.41 | 0.38 | 0.36 | 0.4 | 0.419 | 0.39 | 1
| 1 | 2 | 0.38 | 0.34 | 0.38 | 0.36 | 0.38 | 0.31 | 0.36 | 0.37 | 0.4 | 0.4092 | 0.38 | 0.3903 | 0.33187 | 0.254 | 0.38328 | 0.4 | 0.378 | 0.41 | 0.3824 | 0.3994 | 0.38 | 0.40264 | 0.388 | 0.4 | 0.38 | 0.36 | 0.39 | 0.429 | 0.4 | 0
| 1 | 3 | 0.39 | 0.33 | 0.37 | 0.36 | 0.36 | 0.32 | 0.36 | 0.35 | 0.38 | 0.4043 | 0.36 | 0.3863 | 0.3342 | 0.209 | 0.39224 | 0.41 | 0.386 | 0.41 | 0.3737 | 0.4072 | 0.37 | 0.41542 | 0.391 | 0.4 | 0.38 | 0.36 | 0.39 | 0.424 | 0.41 | 16
| 1 | 4 | 0.4 | 0.33 | 0.38 | 0.36 | 0.36 | 0.33 | 0.35 | 0.35 | 0.4 | 0.3934 | 0.38 | 0.3844 | 0.34708 | 0.301 | 0.39142 | 0.41 | 0.383 | 0.4 | 0.3826 | 0.3976 | 0.36 | 0.40989 | 0.39 | 0.39 | 0.37 | 0.37 | 0.39 | 0.427 | 0.38 | 0.38158
| 1 | 5 | 0.4 | 0.34 | 0.37 | 0.36 | 0.34 | 0.33 | 0.35 | 0.35 | 0.4 | 0.3934 | 0.38 | 0.3892 | 0.32872 | 0.247 | 0.38244 | 0.42 | 0.383 | 0.4 | 0.3835 | 0.39 | 0.38 | 0.4074 | 0.392 | 0.4 | 0.38 | 0.37 | 0.39 | 0.429 | 0.4 | 36
| 1 | 6 | 0.39 | 0.34 | 0.35 | 0.37 | 0.35 | 0.32 | 0.35 | 0.34 | 0.43 | 0.3917 | 0.39 | 0.3864 | 0.3275 | 0.263 | 0.38501 | 0.41 | 0.382 | 0.39 | 0.3521 | 0.3899 | 0.38 | 0.41626 | 0.384 | 0.4 | 0.37 | 0.37 | 0.39 | 0.42 | 0.39
| Mean |  | 0.39167 | 0.33833 | 0.37 | 0.36167 | 0.36 | 0.32167 | 0.355 | 0.35 | 0.40167 | 0.39997 | 0.37667 | 0.39063 | 0.3326 | 0.261 | 0.38829 | 0.41 | 0.38433 | 0.40167 | 0.37647 | 0.3972 | 0.37167 | 0.41159 | 0.38717 | 0.4 | 0.37667 | 0.365 | 0.39167 | 0.42467 | 0.395
| Median |  | 0.39 | 0.34 | 0.37 | 0.36 | 0.36 | 0.32 | 0.355 | 0.35 | 0.4 | 0.39885 | 0.38 | 0.3878 | 0.3303 | 0.2585 | 0.38821 | 0.41 | 0.383 | 0.4 | 0.3825 | 0.39835 | 0.375 | 0.41265 | 0.389 | 0.4 | 0.38 | 0.365 | 0.39 | 0.4255 | 0.395
| Std Dev. |  | 0.0075277 | 0.0075277 | 0.010954 | 0.0040825 | 0.014142 | 0.0075277 | 0.0054772 | 0.010954 | 0.016021 | 0.0079997 | 0.010328 | 0.0083925 | 0.0076683 | 0.033208 | 0.0053871 | 0.0063246 | 0.0053914 | 0.0075277 | 0.012556 | 0.0065382 | 0.0098319 | 0.0059535 | 0.0053072 | 0.0063246 | 0.005164 | 0.0054772 | 0.0040825 | 0.0044121 | 0.010488
| Rel.Std.Dev. |  | 1.922% | 2.225% | 2.961% | 1.129% | 3.928% | 2.34% | 1.543% | 3.13% | 3.989% | 2% | 2.742% | 2.148% | 2.306% | 12.72% | 1.387% | 1.543% | 1.403% | 1.874% | 3.335% | 1.646% | 2.645% | 1.446% | 1.371% | 1.581% | 1.371% | 1.501% | 1.042% | 1.039% | 2.655%
| PDM3 |  | 2.644% | -11.33% | -3.034% | -5.218% | -5.655% | -15.7% | -6.965% | -8.275% | 5.265% | 4.819% | -1.287% | 2.373% | -12.83% | -31.6% | 1.758% | 7.449% | 0.7223% | 5.265% | -1.339% | 4.094% | -2.597% | 7.866% | 1.465% | 4.828% | -1.287% | -4.344% | 2.644% | 11.29% | 3.518%
| Z-Score (Absolute) |  | 0.32 | 1.98 | 0.62 | 0.98 | 1.05 | 2.7 | 1.26 | 1.48 | 0.75 | 0.67 | 0.33 | 0.27 | 2.23 | 5.31 | 0.17 | 1.1 | 0 | 0.75 | 0.34 | 0.55 | 0.55 | 1.17 | 0.12 | 0.67 | 0.33 | 0.83 | 0.32 | 1.74 | 0.46
| Analytical results for Mn in OREAS 627 (Certified Value 0.703 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 3
3
| 1 | 1 | 0.647 | 0.668 | 0.5819 | 0.749 | 0.698 | 0.727 | 0.694 | 0.741 | 0.719 | 0.709 | 0.7171 | 0.7385 | 0.67109 | 0.519 | 0.67644 | 0.6826 | 0.703 | 0.7403 | 0.7509 | 0.6412 | 0.6945 | 0.7328 | 0.7027 | 0.6952 | 0.6862 | 0.6771 | 0.6907 | 0.748 | 0.6997 | 1
| 1 | 2 | 0.6426 | 0.676 | 0.589 | 0.747 | 0.71 | 0.726 | 0.709 | 0.761 | 0.684 | 0.7109 | 0.7482 | 0.7133 | 0.67012 | 0.471 | 0.67327 | 0.6685 | 0.708 | 0.789 | 0.7856 | 0.6427 | 0.7296 | 0.73545 | 0.7014 | 0.7009 | 0.6874 | 0.6813 | 0.6884 | 0.7532 | 0.7071 | 21
| 1 | 3 | 0.6528 | 0.697 | 0.5782 | 0.753 | 0.691 | 0.728 | 0.681 | 0.738 | 0.73 | 0.7002 | 0.7012 | 0.7282 | 0.66713 | 0.571 | 0.67028 | 0.6796 | 0.702 | 0.7441 | 0.8012 | 0.6507 | 0.7074 | 0.73283 | 0.7008 | 0.6936 | 0.6805 | 0.6881 | 0.6683 | 0.7422 | 0.7242 | 16
| 1 | 4 | 0.6608 | 0.694 | 0.5862 | 0.726 | 0.688 | 0.736 | 0.681 | 0.742 | 0.723 | 0.7009 | 0.7302 | 0.7237 | 0.67446 | 0.54 | 0.67239 | 0.6694 | 0.701 | 0.7909 | 0.7953 | 0.6545 | 0.7064 | 0.73941 | 0.7029 | 0.6994 | 0.6802 | 0.6916 | 0.6996 | 0.7445 | 0.6985 | 0.70342
| 1 | 5 | 0.6538 | 0.684 | 0.5824 | 0.744 | 0.69 | 0.732 | 0.685 | 0.74 | 0.702 | 0.7002 | 0.7166 | 0.7292 | 0.68124 | 0.512 | 0.67537 | 0.686 | 0.696 | 0.7885 | 0.7974 | 0.644 | 0.7325 | 0.7394 | 0.711 | 0.6872 | 0.6893 | 0.6683 | 0.6788 | 0.7518 | 0.7215 | 37
| 1 | 6 | 0.6524 | 0.697 | 0.5775 | 0.733 | 0.697 | 0.725 | 0.679 | 0.727 | 0.696 | 0.6977 | 0.7308 | 0.7235 | 0.6788 | 0.548 | 0.67153 | 0.6763 | 0.693 | 0.7247 | 0.7868 | 0.6378 | 0.7108 | 0.74215 | 0.7097 | 0.6849 | 0.6849 | 0.7198 | 0.6785 | 0.7525 | 0.7109
| Mean |  | 0.65157 | 0.686 | 0.58253 | 0.742 | 0.69567 | 0.729 | 0.68817 | 0.7415 | 0.709 | 0.70315 | 0.72402 | 0.72607 | 0.67381 | 0.52683 | 0.67322 | 0.67707 | 0.7005 | 0.76292 | 0.7862 | 0.64515 | 0.71353 | 0.73701 | 0.70475 | 0.69353 | 0.68475 | 0.6877 | 0.68405 | 0.7487 | 0.71032
| Median |  | 0.6526 | 0.689 | 0.58215 | 0.7455 | 0.694 | 0.7275 | 0.683 | 0.7405 | 0.7105 | 0.70055 | 0.72365 | 0.72595 | 0.67278 | 0.5295 | 0.67283 | 0.67795 | 0.7015 | 0.7663 | 0.79105 | 0.64335 | 0.7091 | 0.73742 | 0.7028 | 0.6944 | 0.68555 | 0.6847 | 0.6836 | 0.7499 | 0.709
| Std Dev. |  | 0.0062244 | 0.012116 | 0.0044675 | 0.010354 | 0.0080664 | 0.0041952 | 0.011531 | 0.011005 | 0.017776 | 0.0054121 | 0.016078 | 0.0082962 | 0.0054058 | 0.034557 | 0.0023301 | 0.0070653 | 0.0053198 | 0.029813 | 0.018331 | 0.0062478 | 0.01467 | 0.0038864 | 0.0044275 | 0.0064205 | 0.0037045 | 0.017751 | 0.011058 | 0.0045756 | 0.010778
| Rel.Std.Dev. |  | 0.9553% | 1.766% | 0.7669% | 1.395% | 1.16% | 0.5755% | 1.676% | 1.484% | 2.507% | 0.7697% | 2.221% | 1.143% | 0.8023% | 6.559% | 0.3461% | 1.044% | 0.7594% | 3.908% | 2.332% | 0.9684% | 2.056% | 0.5273% | 0.6282% | 0.9258% | 0.541% | 2.581% | 1.617% | 0.6111% | 1.517%
| PDM3 |  | -7.372% | -2.477% | -17.19% | 5.484% | -1.102% | 3.636% | -2.169% | 5.413% | 0.793% | -0.03862% | 2.928% | 3.219% | -4.21% | -25.1% | -4.294% | -3.747% | -0.4154% | 8.458% | 11.77% | -8.284% | 1.437% | 4.775% | 0.1888% | -1.406% | -2.654% | -2.235% | -2.754% | 6.437% | 0.9802%
| Z-Score (Absolute) |  | 1.4 | 0.42 | 3.38 | 1.19 | 0.14 | 0.82 | 0.35 | 1.18 | 0.24 | 0.08 | 0.67 | 0.73 | 0.77 | 4.98 | 0.78 | 0.67 | 0 | 1.79 | 2.46 | 1.59 | 0.37 | 1.05 | 0.12 | 0.2 | 0.45 | 0.37 | 0.47 | 1.38 | 0.28
| Analytical results for Mo in OREAS 627 (Certified Value 31.3 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 31 | 30.5 | 28 | 31.7 | 32.2 | 32.9 | 28.5 | 31.5 | 31.7 | 31.7 | 29.2 | 31.8 | 27.8774 | 25.699 | 29.1341 | 26.26 | 31.1 | 39.6 | 28.4 | 33.2 | 31.18 | 30.81 | 33 | 29.9 | 32.1 | 33.4 | 32 | 34.76 |  | 1
| 1 | 2 | 31 | 31.9 | 29.1 | 31.2 | 34.1 | 33.3 | 29.3 | 31.6 | 31.8 | 31.2 | 30.2 | 32.6 | 27.4243 | 25.99 | 28.533 | 27.29 | 31.4 | 40.2 | 28.26 | 33.1 | 32.33 | 30.66 | 32 | 32.1 | 31.86 | 33.45 | 33 | 34.71 |  | 37
| 1 | 3 | 32 | 31.7 | 27.6 | 31 | 33.6 | 31.6 | 28.7 | 32.2 | 31.3 | 31.6 | 28.3 | 32.4 | 27.2803 | 26.184 | 28.3788 | 27.61 | 31.5 | 40.3 | 28.14 | 32.7 | 31.48 | 30.39 | 32 | 31.2 | 32.27 | 33.24 | 31 | 31.44 |  | 16
| 1 | 4 | 33 | 31.1 | 28.8 | 30.9 | 33.1 | 33 | 28.6 | 31.9 | 31 | 31 | 29.5 | 33.1 | 27.8218 | 24.783 | 29.1741 | 27.05 | 31.1 | 39.6 | 28.33 | 33.6 | 31.01 | 31.46 | 32 | 29.8 | 32.35 | 33.16 | 32 | 32.4 |  | 31.3337
| 1 | 5 | 32 | 31.2 | 29.3 | 31.4 | 30.5 | 31.6 | 27.5 | 30.9 | 33 | 30.8 | 29.2 | 32.3 | 27.898 | 25.459 | 28.985 | 26.06 | 31.8 | 39.6 | 28.37 | 33.5 | 32.15 | 31.83 | 33 | 31.2 | 32.49 | 33.04 | 32 | 33.61 |  | 38
| 1 | 6 | 32 | 31.3 | 28 | 30.6 | 33.2 | 32.9 | 28.8 | 32.5 | 32.1 | 30.8 | 30.2 | 32.1 | 27.7427 | 25.521 | 28.1521 | 27.02 | 31.7 | 39.5 | 28.36 | 33.6 | 31.99 | 30.17 | 31 | 32 | 32.32 | 32.41 | 32 | 32.65
| Mean |  | 31.8333 | 31.2833 | 28.4667 | 31.1333 | 32.7833 | 32.55 | 28.5667 | 31.7667 | 31.8167 | 31.1833 | 29.4333 | 32.3833 | 27.6741 | 25.606 | 28.7262 | 26.8817 | 31.4333 | 39.8 | 28.31 | 33.2833 | 31.69 | 30.8867 | 32.1667 | 31.0333 | 32.2317 | 33.1167 | 32 | 33.2617
| Median |  | 32 | 31.25 | 28.4 | 31.1 | 33.15 | 32.9 | 28.65 | 31.75 | 31.75 | 31.1 | 29.35 | 32.35 | 27.7823 | 25.61 | 28.759 | 27.035 | 31.45 | 39.6 | 28.345 | 33.35 | 31.735 | 30.735 | 32 | 31.2 | 32.295 | 33.2 | 32 | 33.13
| Std Dev. |  | 0.75277 | 0.4916 | 0.69186 | 0.38816 | 1.2828 | 0.75033 | 0.59217 | 0.56451 | 0.6969 | 0.392 | 0.7174 | 0.4446 | 0.25899 | 0.48915 | 0.42933 | 0.60105 | 0.29439 | 0.35214 | 0.095917 | 0.35449 | 0.54365 | 0.63814 | 0.75277 | 0.99331 | 0.22158 | 0.37782 | 0.63246 | 1.3341
| Rel.Std.Dev. |  | 2.365% | 1.571% | 2.43% | 1.247% | 3.913% | 2.305% | 2.073% | 1.777% | 2.19% | 1.257% | 2.437% | 1.373% | 0.9359% | 1.91% | 1.495% | 2.236% | 0.9366% | 0.8848% | 0.3388% | 1.065% | 1.716% | 2.066% | 2.34% | 3.201% | 0.6875% | 1.141% | 1.976% | 4.011%
| PDM3 |  | 1.594% | -0.1608% | -9.15% | -0.6396% | 4.626% | 3.882% | -8.831% | 1.382% | 1.541% | -0.48% | -6.065% | 3.35% | -11.68% | -18.28% | -8.322% | -14.21% | 0.3179% | 27.02% | -9.65% | 6.222% | 1.137% | -1.427% | 2.658% | -0.9587% | 2.866% | 5.69% | 2.126% | 6.153%
| Z-Score (Absolute) |  | 0.2 | 0.21 | 2.31 | 0.32 | 0.91 | 0.74 | 2.23 | 0.15 | 0.19 | 0.28 | 1.59 | 0.61 | 2.9 | 4.44 | 2.11 | 3.49 | 0.1 | 6.14 | 2.42 | 1.28 | 0.1 | 0.5 | 0.45 | 0.39 | 0.5 | 1.16 | 0.33 | 1.27
| Analytical results for Na in OREAS 627 (Certified Value 1.07 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.05 | 0.96 | 1.074 | 1.12 | 1.02 | 1.08 | 1.07 | 1.11 | 1.07 | 1.049 | 1.066 | 1.0837 | 1.0029 | 1.0489 | 1.05 | 1.07 | 1.09 | 1.0167 | 0.8341 | 1.29 | 1.0258 | 1.046 | 1.13 | 1.03 | 1.21 | 1.11 | 1.125 | 1.11 |  | 1
| 1 | 2 | 1.05 | 1.01 | 1.07 | 1.13 | 1.05 | 1.06 | 1.11 | 1.16 | 1.05 | 1.055 | 1.11 | 1.1005 | 0.98906 | 1.0414 | 1.04 | 1.09 | 1.06 | 1.0754 | 0.8447 | 1.4 | 1.0262 | 1.037 | 1.1 | 1.04 | 1.22 | 1.13 | 1.093 | 1.11 |  | 13
| 1 | 3 | 1.04 | 1 | 1.058 | 1.12 | 1.01 | 1.07 | 1.08 | 1.12 | 1.03 | 1.0265 | 1.038 | 1.0873 | 0.9848 | 1.058 | 1.08 | 1.08 | 1.07 | 1.0845 | 0.8688 | 1.35 | 1.062 | 1.026 | 1.1 | 1.04 | 1.19 | 1.1 | 1.095 | 1.13 |  | 16
| 1 | 4 | 1.06 | 1.01 | 1.06 | 1.11 | 1.04 | 1.06 | 1.1 | 1.13 | 1.04 | 1.0494 | 1.059 | 1.0864 | 0.98918 | 1.061 | 1.03 | 1.07 | 1.06 | 1.0832 | 0.9017 | 1.25 | 1.0143 | 1.036 | 1.09 | 1.04 | 1.19 | 1.12 | 1.101 | 1.12 |  | 1.0671
| 1 | 5 | 1.03 | 1.01 | 1.07 | 1.1 | 0.99 | 1.04 | 1.07 | 1.14 | 1.06 | 1.0568 | 1.046 | 1.0951 | 1.0105 | 1.0442 | 1.04 | 1.07 | 1.07 | 1.0776 | 0.8728 | 1.37 | 1.0265 | 1.036 | 1.11 | 1.04 | 1.2 | 1.1 | 1.107 | 1.15 |  | 39
| 1 | 6 | 1.07 | 1.03 | 1.066 | 1.08 | 1.06 | 1.05 | 1.06 | 1.14 | 1.09 | 1.0598 | 1.063 | 1.0861 | 1.035 | 1.0615 | 1.05 | 1.09 | 1.09 | 1.0812 | 0.8723 | 1.33 | 1.0472 | 1.064 | 1.05 | 1.04 | 1.19 | 1.12 | 1.115 | 1.14
| Mean |  | 1.05 | 1.0033 | 1.0663 | 1.11 | 1.0283 | 1.06 | 1.0817 | 1.1333 | 1.0567 | 1.0494 | 1.0637 | 1.0898 | 1.0019 | 1.0525 | 1.0483 | 1.0783 | 1.0733 | 1.0698 | 0.86573 | 1.3317 | 1.0337 | 1.0408 | 1.0967 | 1.0383 | 1.2 | 1.1133 | 1.106 | 1.1267
| Median |  | 1.05 | 1.01 | 1.068 | 1.115 | 1.03 | 1.06 | 1.075 | 1.135 | 1.055 | 1.0522 | 1.061 | 1.0869 | 0.99602 | 1.0534 | 1.045 | 1.075 | 1.07 | 1.0794 | 0.87055 | 1.34 | 1.0264 | 1.0365 | 1.1 | 1.04 | 1.195 | 1.115 | 1.104 | 1.125
| Std Dev. |  | 0.014142 | 0.023381 | 0.0062503 | 0.017889 | 0.026394 | 0.014142 | 0.019408 | 0.017512 | 0.021602 | 0.01199 | 0.025081 | 0.0065004 | 0.01893 | 0.0088203 | 0.017224 | 0.0098319 | 0.013663 | 0.02622 | 0.023829 | 0.054559 | 0.017493 | 0.012999 | 0.026583 | 0.0040825 | 0.012649 | 0.012111 | 0.012313 | 0.01633
| Rel.Std.Dev. |  | 1.347% | 2.33% | 0.5862% | 1.612% | 2.567% | 1.334% | 1.794% | 1.545% | 2.044% | 1.143% | 2.358% | 0.5964% | 1.889% | 0.838% | 1.643% | 0.9118% | 1.273% | 2.451% | 2.752% | 4.097% | 1.692% | 1.249% | 2.424% | 0.3932% | 1.054% | 1.088% | 1.113% | 1.449%
| PDM3 |  | -1.599% | -5.973% | -0.06854% | 4.024% | -3.63% | -0.6621% | 1.368% | 6.21% | -0.9745% | -1.654% | -0.3184% | 2.135% | -6.107% | -1.365% | -1.755% | 1.056% | 0.5875% | 0.2532% | -18.87% | 24.8% | -3.13% | -2.458% | 2.774% | -2.693% | 12.46% | 4.336% | 3.649% | 5.586%
| Z-Score (Absolute) |  | 0.39 | 1.61 | 0.03 | 1.18 | 0.96 | 0.13 | 0.44 | 1.79 | 0.22 | 0.41 | 0.03 | 0.65 | 1.65 | 0.33 | 0.44 | 0.35 | 0.22 | 0.12 | 5.22 | 6.98 | 0.82 | 0.63 | 0.83 | 0.7 | 3.53 | 1.27 | 1.07 | 1.61
| Analytical results for Nb in OREAS 627 (Certified Value 10.4 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 1
2
| 1 | 1 | 6 | 10.3 | 10.2 | 10.8 | 12.1 | 11.9 | 9.8 | 11.2 | 10.1 | 10.28 | 8.3 | 10.55 | 9.3372 | 9.8007 | 9.6 | 10.3 | 9.33 | 10.5 | 9.4 | 8 | 9.3 | 9.1 | 11.2 | 13 | [5 empty] | 1
| 1 | 2 | 6 | 10.6 | 10.5 | 10.6 | 12.5 | 12.5 | 10.1 | 11.2 | 10.5 | 10.53 | 8.3 | 11.15 | 9.2652 | 9.6623 | 9.9 | 11.1 | 9.38 | 11.1 | 9.1 | 8 | 9.4 | 8.3 | 11.5 | 12 | [5 empty] | 38
| 1 | 3 | 6 | 10.8 | 9.6 | 10.5 | 12.3 | 11.1 | 10 | 10.9 | 10 | 10.82 | 8 | 11.13 | 9.4209 | 9.6829 | 9.7 | 10.8 | 9.06 | 10.5 | 9.1 | 9 | 9.7 | 9.1 | 11.2 | 11.1 | [5 empty] | 16
| 1 | 4 | 6 | 9.7 | 10.2 | 10.6 | 12.4 | 11.8 | 9.9 | 11.2 | 10.2 | 10.91 | 8.5 | 11.19 | 9.4592 | 9.6411 | 9.9 | 10.8 | 9.37 | 10.5 | 9.8 | 8 | 9.8 | 8.9 | 11.4 | 12.5 | [5 empty] | 10.374
| 1 | 5 | 6 | 10.7 | 9.9 | 10.4 | 9.8 | 11.7 | 9.8 | 11 | 10.9 | 10.71 | 8.5 | 10.83 | 9.5645 | 9.8077 | 9.9 | 11 | 9.2 | 10.9 | 9.6 | 9 | 9.6 | 9.3 | 11.2 | 13.8 | [5 empty] | 40
| 1 | 6 | 6 | 10.6 | 10.2 | 10.4 | 12.2 | 11.5 | 10.5 | 11.6 | 10.4 | 10.59 | 8.8 | 10.98 | 9.5974 | 9.6381 | 9.8 | 10.7 | 9.14 | 10.9 | 9.8 | 8 | 9.9 | 9.6 | 11.2 | 12.6
| Mean |  | 6 | 10.45 | 10.1 | 10.55 | 11.8833 | 11.75 | 10.0167 | 11.1833 | 10.35 | 10.64 | 8.4 | 10.9717 | 9.4407 | 9.7055 | 9.8 | 10.7833 | 9.2467 | 10.7333 | 9.4667 | 8.3333 | 9.6167 | 9.05 | 11.2833 | 12.5
| Median |  | 6 | 10.6 | 10.2 | 10.55 | 12.25 | 11.75 | 9.95 | 11.2 | 10.3 | 10.65 | 8.4 | 11.055 | 9.4401 | 9.6726 | 9.85 | 10.8 | 9.265 | 10.7 | 9.5 | 8 | 9.65 | 9.1 | 11.2 | 12.55
| Std Dev. |  | 0 | 0.40373 | 0.30984 | 0.15166 | 1.0304 | 0.46368 | 0.26394 | 0.24014 | 0.32711 | 0.22557 | 0.26833 | 0.24596 | 0.12818 | 0.078192 | 0.12649 | 0.27869 | 0.13292 | 0.26583 | 0.32042 | 0.5164 | 0.23166 | 0.43704 | 0.13292 | 0.91214
| Rel.Std.Dev. |  | 0% | 3.863% | 3.068% | 1.438% | 8.671% | 3.946% | 2.635% | 2.147% | 3.16% | 2.12% | 3.194% | 2.242% | 1.358% | 0.8056% | 1.291% | 2.584% | 1.437% | 2.477% | 3.385% | 6.197% | 2.409% | 4.829% | 1.178% | 7.297%
| PDM3 |  | -42.16% | 0.7327% | -2.641% | 1.697% | 14.55% | 13.26% | -3.444% | 7.802% | -0.2313% | 2.564% | -19.03% | 5.761% | -8.996% | -6.444% | -5.533% | 3.946% | -10.87% | 3.464% | -8.746% | -19.67% | -7.3% | -12.76% | 8.766% | 20.49%
| Z-Score (Absolute) |  | NA | 0.05 | 0.27 | 0.14 | 1.35 | 1.23 | 0.35 | 0.71 | 0.05 | 0.22 | 1.83 | 0.52 | 0.88 | 0.63 | 0.55 | 0.35 | 1.05 | 0.3 | 0.85 | NA | 0.71 | 1.23 | 0.81 | 1.92
| Analytical results for Nd in OREAS 627 (Certified Value 31.4 ppm) | [30 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | [5 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | [5 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | [5 empty] | 1
2
| 1 | 1 | 1.01: 25.8; 1.02: 31.99; 1.03: 31.48; 1.04: 20.0203; 1.05: 34.1509; 1.06: 31.1 | [5 empty] | 1
| 1 | 2 | 1.01: 25.7; 1.02: 31.42; 1.03: 33.63; 1.04: 19.5876; 1.05: 34.7505; 1.06: 33.6 | [5 empty] | 7
| 1 | 3 | 1.01: 27.8; 1.02: 32.03; 1.03: 32.57; 1.04: 20.6677; 1.05: 34.6797; 1.06: 31.7 | [5 empty] | 16
| 1 | 4 | 1.01: 26.1; 1.02: 30.58; 1.03: 33.34; 1.04: 21.6955; 1.05: 34.9828; 1.06: 31 | [5 empty] | 31.3774
| 1 | 5 | 1.01: 26.7; 1.02: 29.84; 1.03: 32.56; 1.04: 19.3898; 1.05: 34.5395; 1.06: 32.7 | [5 empty] | 41
| 1 | 6 | 1.01: 26.5; 1.02: 30.72; 1.03: 33.24; 1.04: 20.2021; 1.05: 34.7192; 1.06: 31.4
| Mean |  | 1.01: 26.4333; 1.02: 31.0967; 1.03: 32.8033; 1.04: 20.2605; 1.05: 34.6371; 1.06: 31.9167
| Median |  | 1.01: 26.3; 1.02: 31.07; 1.03: 32.905; 1.04: 20.1112; 1.05: 34.6995; 1.06: 31.55
| Std Dev. |  | 1.01: 0.77374; 1.02: 0.86738; 1.03: 0.77807; 1.04: 0.8363; 1.05: 0.27815; 1.06: 1.0265
| Rel.Std.Dev. |  | 1.01: 2.927%; 1.02: 2.789%; 1.03: 2.372%; 1.04: 4.128%; 1.05: 0.803%; 1.06: 3.216%
| PDM3 |  | 1.01: -15.76%; 1.02: -0.8948%; 1.03: 4.544%; 1.04: -35.43%; 1.05: 10.39%; 1.06: 1.719%
| Z-Score (Absolute) |  | 1.01: 1.55; 1.02: 0.12; 1.03: 0.4; 1.04: 3.43; 1.05: 0.95; 1.06: 0.12
| Analytical results for Ni in OREAS 627 (Certified Value 7.69 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 7 | 7 | 6.6 | 7.3 | 8.3 | 7.2 | 7.6 | 8.3 | 8.3 | 7.8 | 7.7 | 8.1 | 6.6961 | 5.084 | 7.4306 | 9.3 | 7 | 7.5 | 6.68 | 8.4 | 7.7 | 10 | 7 | 7.9 | 7.4 | 7.4 | 6 | 8.8 |  | 1
| 1 | 2 | 8 | 7.1 | 5.8 | 7 | 8.6 | 7.3 | 7.4 | 8.2 | 7.9 | 8 | 7.9 | 7.9 | 7.221 | 4.868 | 7.6309 | 8.55 | 7 | 7.9 | 6.85 | 8.2 | 8 | 10.48 | 8 | 8.1 | 7.8 | 7.3 | 7 | 13 |  | 14
| 1 | 3 | 8 | 6.9 | 5.8 | 7.5 | 8.5 | 6.8 | 7.4 | 8.1 | 7.6 | 8.5 | 7.4 | 7.8 | 7.257 | 4.65 | 7.6051 | 8.24 | 7 | 7.5 | 7.05 | 7.7 | 7.6 | 10.14 | 8 | 8.3 | 7.7 | 7.2 | 6 | 7.7 |  | 16
| 1 | 4 | 8 | 7.1 | 5.4 | 7.5 | 8 | 7.6 | 7.2 | 8.3 | 7.7 | 7.7 | 7.9 | 8 | 7.0164 | 4.841 | 7.4172 | 7.59 | 7 | 7.9 | 6.83 | 8 | 8.3 | 9.61 | 8 | 7.6 | 8 | 7.4 | 6 | 7.8 |  | 7.6919
| 1 | 5 | 9 | 7.1 | 5.6 | 7 | 7.1 | 7.6 | 7 | 8 | 7.7 | 7.4 | 9 | 7.9 | 7.3723 | 5.324 | 7.3861 | 8.9 | 7 | 7.4 | 6.89 | 8.7 | 8.1 | 9.68 | 8 | 7.8 | 7.9 | 7.5 | 7 | 8.6 |  | 42
| 1 | 6 | 8 | 7.2 | 5.4 | 7.5 | 8 | 7.8 | 7.2 | 8.2 | 8.6 | 7.9 | 8.2 | 8.2 | 6.7382 | 5.038 | 7.3309 | 8.5 | 7 | 7.5 | 6.88 | 7 | 8.2 | 9.83 | 7 | 7.8 | 8.5 | 7.5 | 6 | 7.3
| Mean |  | 8 | 7.0667 | 5.7667 | 7.3 | 8.0833 | 7.3833 | 7.3 | 8.1833 | 7.9667 | 7.8833 | 8.0167 | 7.9833 | 7.0502 | 4.9675 | 7.4668 | 8.5133 | 7 | 7.6167 | 6.8633 | 8 | 7.9833 | 9.9567 | 7.6667 | 7.9167 | 7.8833 | 7.3833 | 6.3333 | 8.8667
| Median |  | 8 | 7.1 | 5.7 | 7.4 | 8.15 | 7.45 | 7.3 | 8.2 | 7.8 | 7.85 | 7.9 | 7.95 | 7.1187 | 4.953 | 7.4239 | 8.525 | 7 | 7.5 | 6.865 | 8.1 | 8.05 | 9.915 | 8 | 7.85 | 7.85 | 7.4 | 6 | 8.2
| Std Dev. |  | 0.63246 | 0.10328 | 0.44572 | 0.24495 | 0.54191 | 0.36009 | 0.20976 | 0.1169 | 0.39833 | 0.3656 | 0.54924 | 0.1472 | 0.2827 | 0.23339 | 0.12232 | 0.5826 | 0 | 0.22286 | 0.11894 | 0.59666 | 0.27869 | 0.32315 | 0.5164 | 0.24833 | 0.3656 | 0.1169 | 0.5164 | 2.103
| Rel.Std.Dev. |  | 7.906% | 1.462% | 7.729% | 3.355% | 6.704% | 4.877% | 2.873% | 1.429% | 5% | 4.638% | 6.851% | 1.844% | 4.01% | 4.698% | 1.638% | 6.843% | 0% | 2.926% | 1.733% | 7.458% | 3.491% | 3.246% | 6.736% | 3.137% | 4.638% | 1.583% | 8.154% | 23.72%
| PDM3 |  | 4.006% | -8.128% | -25.03% | -5.095% | 5.089% | -4.011% | -5.095% | 6.389% | 3.572% | 2.489% | 4.222% | 3.789% | -8.343% | -35.42% | -2.926% | 10.68% | -8.995% | -0.9779% | -10.77% | 4.006% | 3.789% | 29.44% | -0.3279% | 2.922% | 2.489% | -4.011% | -17.66% | 15.27%
| Z-Score (Absolute) |  | NA | 1.2 | 3.11 | 0.86 | 0.29 | 0.74 | 0.86 | 0.44 | 0.12 | 0 | 0.2 | 0.15 | 1.23 | 4.29 | 0.61 | 0.93 | NA | 0.39 | 1.5 | 0.17 | 0.15 | 3.05 | NA | 0.05 | 0 | 0.74 | NA | 1.45
Results from laboratories 1.01, 1.17, 1.23 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 627 (Certified Value 0.037 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.0411 | 0.038 | 0.038 | 0.036 | 0.036 | 0.038 | 0.038 | 0.0372 | 0.036 | 0.035 | 0.0373 | 0.03456 | 0.03479 | 0.04 | 0.036 | 0.037 | 0.0358 | 0.0379 | 0.04 | 0.036 | 0.034 | 0.04 | 0.0336 | 0.0381 | 0.03 | 0.0365 |  |  |  | 1
| 1 | 2 | 0.0403 | 0.039 | 0.039 | 0.036 | 0.035 | 0.038 | 0.041 | 0.0372 | 0.0358 | 0.037 | 0.0367 | 0.035078 | 0.034663 | 0.04 | 0.037 | 0.037 | 0.036 | 0.0381 | 0.04 | 0.037 | 0.034 | 0.04 | 0.0339 | 0.0381 | 0.033 | 0.0374 |  |  |  | 24
| 1 | 3 | 0.0396 | 0.037 | 0.038 | 0.035 | 0.035 | 0.037 | 0.04 | 0.0356 | 0.0353 | 0.035 | 0.0361 | 0.034285 | 0.036253 | 0.04 | 0.036 | 0.038 | 0.0364 | 0.0392 | 0.04 | 0.038 | 0.034 | 0.04 | 0.034 | 0.0364 | 0.031 | 0.0369 |  |  |  | 16
| 1 | 4 | 0.0415 | 0.039 | 0.037 | 0.035 | 0.037 | 0.036 | 0.04 | 0.0366 | 0.0346 | 0.037 | 0.0357 | 0.035489 | 0.036398 | 0.04 | 0.036 | 0.038 | 0.0362 | 0.0378 | 0.04 | 0.037 | 0.035 | 0.04 | 0.0347 | 0.0372 | 0.034 | 0.0351 |  |  |  | 0.037017
| 1 | 5 | 0.0406 | 0.038 | 0.038 | 0.036 | 0.036 | 0.036 | 0.04 | 0.0367 | 0.0337 | 0.037 | 0.0357 | 0.03574 | 0.035429 | 0.04 | 0.036 | 0.038 | 0.0363 | 0.0365 | 0.04 | 0.037 | 0.034 | 0.04 | 0.0345 | 0.0366 | 0.032 | 0.0367 |  |  |  | 43
| 1 | 6 | 0.0388 | 0.037 | 0.038 | 0.034 | 0.036 | 0.036 | 0.039 | 0.0399 | 0.0344 | 0.038 | 0.0354 | 0.034716 | 0.035884 | 0.04 | 0.037 | 0.036 | 0.0365 | 0.0366 | 0.04 | 0.036 | 0.035 | 0.04 | 0.0347 | 0.0369 | 0.032 | 0.0361
| Mean |  | 0.040317 | 0.038 | 0.038 | 0.035333 | 0.035833 | 0.036833 | 0.039667 | 0.0372 | 0.034967 | 0.0365 | 0.03615 | 0.034978 | 0.035569 | 0.04 | 0.036333 | 0.037333 | 0.0362 | 0.037683 | 0.04 | 0.036833 | 0.034333 | 0.04 | 0.034233 | 0.037217 | 0.032 | 0.03645
| Median |  | 0.04045 | 0.038 | 0.038 | 0.0355 | 0.036 | 0.0365 | 0.04 | 0.03695 | 0.03495 | 0.037 | 0.0359 | 0.034897 | 0.035656 | 0.04 | 0.036 | 0.0375 | 0.03625 | 0.03785 | 0.04 | 0.037 | 0.034 | 0.04 | 0.03425 | 0.03705 | 0.032 | 0.0366
| Std Dev. |  | 0.00099079 | 0.00089443 | 0.00063246 | 0.0008165 | 0.00075277 | 0.00098319 | 0.0010328 | 0.0014464 | 0.00088694 | 0.0012247 | 0.00072042 | 0.00056123 | 0.00073503 | 0 | 0.0005164 | 0.0008165 | 0.00026077 | 0.0010108 | 0 | 0.00075277 | 0.0005164 | 0 | 0.00046332 | 0.00073598 | 0.0014142 | 0.0007893
| Rel.Std.Dev. |  | 2.458% | 2.354% | 1.664% | 2.311% | 2.101% | 2.669% | 2.604% | 3.888% | 2.537% | 3.355% | 1.993% | 1.605% | 2.066% | 0% | 1.421% | 2.187% | 0.7204% | 2.682% | 0% | 2.044% | 1.504% | 0% | 1.353% | 1.978% | 4.419% | 2.165%
| PDM3 |  | 8.914% | 2.656% | 2.656% | -4.548% | -3.198% | -0.4961% | 7.158% | 0.4945% | -5.539% | -1.397% | -2.342% | -5.508% | -3.91% | 8.059% | -1.847% | 0.8546% | -2.207% | 1.8% | 8.059% | -0.4961% | -7.25% | 8.059% | -7.52% | 0.5395% | -13.55% | -1.532%
| Z-Score (Absolute) |  | 2.03 | 0.74 | 0.74 | 0.74 | 0.46 | 0.09 | 1.66 | 0.3 | 0.94 | 0.09 | 0.29 | 0.94 | 0.61 | 1.85 | 0.18 | 0.37 | 0.26 | 0.56 | 1.85 | 0.09 | 1.29 | 1.85 | 1.35 | 0.31 | 2.59 | 0.12
| Analytical results for Pb in OREAS 627 (Certified Value 2.26 wt.%) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES/MS | 2
3
| 1 | 1 | 2.42 | 2.22 | 2.236 | 2.18 | 2.21 | 2.24 | 2.23 | 2.25 | 2.25 | 2.2934 | 2.4 | 2.1992 | 2.2328 | 2.1468 | 2.2939 | 2.3 | 2.19 | 1.82 | 2.3281 | 2.3535 | 2.326 | 2.1892 | 2.19 | 2.32 | 2.33 | 2.244 | 2.2 | > 1 | 2.1804 | 1
| 1 | 2 | 2.45 | 2.2 | 2.31 | 2.21 | 2.18 | 2.22 | 2.24 | 2.24 | 2.34 | 2.2496 | 2.38 | 2.3282 | 2.2564 | 2.0542 | 2.2794 | 2.3 | 2.22 | 1.82 | 2.3867 | 2.3663 | 2.321 | 2.157 | 2.21 | 2.25 | 2.3 | 2.236 | 2.21 | > 1 | 2.2423 | 25
| 1 | 3 | 2.3 | 2.21 | 2.21 | 2.22 | 2.21 | 2.22 | 2.22 | 2.23 | 2.54 | 2.245 | 2.36 | 2.1099 | 2.2607 | 1.9004 | 2.2937 | 2.3 | 2.18 | 1.81 | 2.4093 | 2.3665 | 2.278 | 2.2183 | 2.21 | 2.26 | 2.31 | 2.247 | 2.2 | > 1 | 2.331 | 16
| 1 | 4 | 2.33 | 2.21 | 2.257 | 2.22 | 2.19 | 2.24 | 2.22 | 2.22 | 2.38 | 2.2899 | 2.38 | > 1 | 2.2737 | 2.064 | 2.235 | 2.28 | 2.21 | 1.8 | 2.3857 | 2.312 | 2.338 | 2.1424 | 2.21 | 2.28 | 2.35 | 2.23 | 2.2 | > 1 | 2.1696 | 2.2646
| 1 | 5 | 2.39 | 2.19 | 2.238 | 2.23 | 2.23 | 2.27 | 2.21 | 2.24 | 2.26 | 2.2314 | 2.4 | 2.2445 | 2.275 | 2.1311 | 2.2128 | 2.3 | 2.28 | 1.82 | 2.3865 | 2.3598 | 2.318 | 2.1531 | 2.21 | 2.27 | 2.32 | 2.229 | 2.2 | > 1 | 2.2872 | 44
| 1 | 6 | 2.37 | 2.19 | 2.233 | 2.27 | 2.18 | 2.26 | 2.2 | 2.26 | 2.39 | 2.2532 | 2.4 | 2.2932 | 2.2799 | 2.0876 | 2.2248 | 2.22 | 2.21 | 1.81 | 2.3956 | 2.3304 | 2.328 | 2.1998 | 2.19 | 2.23 | 2.32 | 2.243 | 2.22 | > 1 | 2.2512
| Mean |  | 2.3767 | 2.2033 | 2.2473 | 2.2217 | 2.2 | 2.2417 | 2.22 | 2.24 | 2.36 | 2.2604 | 2.3867 | 2.235 | 2.2631 | 2.064 | 2.2566 | 2.2833 | 2.215 | 1.8133 | 2.382 | 2.3481 | 2.3182 | 2.1766 | 2.2033 | 2.2683 | 2.3217 | 2.2382 | 2.205 |  | 2.2436
| Median |  | 2.38 | 2.205 | 2.237 | 2.22 | 2.2 | 2.24 | 2.22 | 2.24 | 2.36 | 2.2514 | 2.39 | 2.2445 | 2.2672 | 2.0758 | 2.2572 | 2.3 | 2.21 | 1.815 | 2.3866 | 2.3567 | 2.3235 | 2.1731 | 2.21 | 2.265 | 2.32 | 2.2395 | 2.2 |  | 2.2468
| Std Dev. |  | 0.055737 | 0.012111 | 0.034162 | 0.029269 | 0.02 | 0.020412 | 0.014142 | 0.014142 | 0.10602 | 0.025321 | 0.01633 | 0.085287 | 0.017355 | 0.088071 | 0.036564 | 0.032042 | 0.035071 | 0.008165 | 0.027898 | 0.022147 | 0.020846 | 0.030139 | 0.010328 | 0.030605 | 0.017224 | 0.0076267 | 0.0083666 |  | 0.061757
| Rel.Std.Dev. |  | 2.345% | 0.5496% | 1.52% | 1.317% | 0.9091% | 0.9106% | 0.637% | 0.6313% | 4.492% | 1.12% | 0.6842% | 3.816% | 0.7669% | 4.267% | 1.62% | 1.403% | 1.583% | 0.4503% | 1.171% | 0.9432% | 0.8993% | 1.385% | 0.4687% | 1.349% | 0.7419% | 0.3408% | 0.3794% |  | 2.753%
| PDM3 |  | 4.948% | -2.706% | -0.7627% | -1.896% | -2.853% | -1.013% | -1.97% | -1.087% | 4.212% | -0.185% | 5.39% | -1.307% | -0.06629% | -8.858% | -0.3541% | 0.827% | -2.19% | -19.93% | 5.183% | 3.686% | 2.365% | -3.885% | -2.706% | 0.1646% | 2.52% | -1.168% | -2.632% |  | -0.9268%
| Z-Score (Absolute) |  | 2.35 | 0.69 | 0.08 | 0.37 | 0.75 | 0.02 | 0.4 | 0.05 | 2.06 | 0.31 | 2.52 | 0.13 | 0.36 | 3.13 | 0.24 | 0.71 | 0.48 | 7.52 | 2.44 | 1.85 | 1.32 | 1.16 | 0.69 | 0.45 | 1.38 | 0.08 | 0.66 | NA | 0.02
| Analytical results for Pr in OREAS 627 (Certified Value 7.53 ppm) | [30 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | 2
2
| 1 | 1 | 1.01: 5.7; 1.02: 7.82; 1.03: 8.14; 1.04: 5.1986; 1.05: 9.7227; 1.06: 8.16 | 1
| 1 | 2 | 1.01: 5.9; 1.02: 7.85; 1.03: 8.58; 1.04: 5.0009; 1.05: 9.0521; 1.06: 9.04 | 9
| 1 | 3 | 1.01: 6.2; 1.02: 8.01; 1.03: 8.41; 1.04: 5.1188; 1.05: 9.5553; 1.06: 8.48 | 16
| 1 | 4 | 1.01: 5.8; 1.02: 7.52; 1.03: 8.54; 1.04: 5.483; 1.05: 9.7414; 1.06: 8.16 | 7.528
| 1 | 5 | 1.01: 5.8; 1.02: 7.49; 1.03: 8.36; 1.04: 5.0098; 1.05: 9.3039; 1.06: 8.73 | 45
| 1 | 6 | 1.01: 5.9; 1.02: 7.63; 1.03: 8.72; 1.04: 5.1845; 1.05: 9.2666; 1.06: 8.43
| Mean |  | 1.01: 5.8833; 1.02: 7.72; 1.03: 8.4583; 1.04: 5.1659; 1.05: 9.4403; 1.06: 8.5
| Median |  | 1.01: 5.85; 1.02: 7.725; 1.03: 8.475; 1.04: 5.1516; 1.05: 9.4296; 1.06: 8.455
| Std Dev. |  | 1.01: 0.17224; 1.02: 0.20591; 1.03: 0.20163; 1.04: 0.17652; 1.05: 0.27683; 1.06: 0.34088
| Rel.Std.Dev. |  | 1.01: 2.928%; 1.02: 2.667%; 1.03: 2.384%; 1.04: 3.417%; 1.05: 2.932%; 1.06: 4.01%
| PDM3 |  | 1.01: -21.85%; 1.02: 2.551%; 1.03: 12.36%; 1.04: -31.38%; 1.05: 25.4%; 1.06: 12.91%
| Z-Score (Absolute) |  | 1.01: 1.69; 1.02: 0.28; 1.03: 0.28; 1.04: 2.24; 1.05: 1.03; 1.06: 0.31
| Analytical results for Rb in OREAS 627 (Certified Value 122 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 0
0
| 1 | 1 | 114 | 113.9 | 123 | 133 | 105 | 107.5 | 117.5 | 140 | 127.2 | 124.6 | 122.21 | 106.6157 | 120.441 | 128 | 136.3 | 125.2 | 113 | 120 | 149 | 118.8 | 128.8 | 137.2 | [7 empty] | 1
| 1 | 2 | 120 | 119.5 | 121.5 | 136 | 106 | 111.5 | 120.5 | 131 | 125.21 | 126.2 | 127.89 | 102.8069 | 119.7523 | 130 | 140.2 | 123.4 | 122 | 116 | 159 | 110.3 | 125.7 | 144 | [7 empty] | 43
| 1 | 3 | 126 | 104.7 | 123.5 | 133 | 105 | 111.5 | 122.5 | 138 | 129.34 | 119.3 | 126.51 | 102.9125 | 119.462 | 129 | 139.2 | 121.9 | 117 | 116 | 167 | 121.9 | 127.4 | 135.1 | [7 empty] | 16
| 1 | 4 | 119 | 115.4 | 125.5 | 133.5 | 107.5 | 108 | 119 | 133 | 124.76 | 121.8 | 129.23 | 101.0848 | 121.0389 | 128 | 135.1 | 126.4 | 112 | 121 | 146 | 121 | 129.1 | 133.2 | [7 empty] | 122.3809
| 1 | 5 | 124 | 108.2 | 124.5 | 115.5 | 105.5 | 107 | 116 | 138 | 121.76 | 120.8 | 127.25 | 98.2152 | 121.389 | 131 | 133.4 | 124 | 120 | 118 | 153 | 122.9 | 126.9 | 140.5 | [7 empty] | 46
| 1 | 6 | 122 | 109.3 | 123.5 | 130.5 | 108 | 111 | 122 | 130 | 126.35 | 123.7 | 129.37 | 105.8895 | 119.4409 | 129 | 125.7 | 124 | 119 | 118 | 164 | 127.5 | 126.6 | 133.1
| Mean |  | 120.8333 | 111.8333 | 123.5833 | 130.25 | 106.1667 | 109.4167 | 119.5833 | 135 | 125.77 | 122.7333 | 127.0767 | 102.9208 | 120.254 | 129.1667 | 134.9833 | 124.15 | 117.1667 | 118.1667 | 156.3333 | 120.4 | 127.4167 | 137.1833
| Median |  | 121 | 111.6 | 123.5 | 133 | 105.75 | 109.5 | 119.75 | 135.5 | 125.78 | 122.75 | 127.57 | 102.8597 | 120.0966 | 129 | 135.7 | 124 | 118 | 118 | 156 | 121.45 | 127.15 | 136.15
| Std Dev. |  | 4.2151 | 5.4088 | 1.3571 | 7.4347 | 1.291 | 2.1311 | 2.5577 | 4.1952 | 2.5511 | 2.5626 | 2.6303 | 3.0981 | 0.83417 | 1.169 | 5.2044 | 1.5385 | 3.9707 | 2.0412 | 8.3825 | 5.7243 | 1.3136 | 4.3439
| Rel.Std.Dev. |  | 3.488% | 4.836% | 1.098% | 5.708% | 1.216% | 1.948% | 2.139% | 3.108% | 2.028% | 2.088% | 2.07% | 3.01% | 0.6937% | 0.9051% | 3.856% | 1.239% | 3.389% | 1.727% | 5.362% | 4.754% | 1.031% | 3.167%
| PDM3 |  | -1.265% | -8.619% | 0.9826% | 6.43% | -13.25% | -10.59% | -2.286% | 10.31% | 2.769% | 0.288% | 3.837% | -15.9% | -1.738% | 5.545% | 10.3% | 1.446% | -4.261% | -3.444% | 27.74% | -1.619% | 4.115% | 12.1%
| Z-Score (Absolute) |  | 0.29 | 1.39 | 0.05 | 0.87 | 2.09 | 1.69 | 0.44 | 1.45 | 0.32 | 0.05 | 0.48 | 2.48 | 0.36 | 0.74 | 1.45 | 0.12 | 0.74 | 0.61 | 4.07 | 0.34 | 0.52 | 1.72
| Analytical results for Re in OREAS 627 (Certified Value < 0.005 ppm) | [30 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [11 empty] | 3
3
| 1 | 1 | < 0.001 | < 0.005 | 0.002 | 0.002 | < 0.002 | 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.005 | < 0.002 | < 0.01 | < 0.05 | < 0.05 | < 0.005 | 0.003 | < 0.002 | < 0.002 | [11 empty] | 1
| 1 | 2 | 0.001 | < 0.005 | < 0.002 | 0.002 | 0.002 | 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.005 | < 0.002 | < 0.01 | < 0.05 | < 0.05 | < 0.005 | 0.004 | 0.002 | < 0.002 | [11 empty] | 15
| 1 | 3 | 0.001 | < 0.005 | < 0.002 | < 0.002 | 0.002 | < 0.002 | 0.002 | 0.002 | < 0.002 | < 0.005 | < 0.002 | < 0.01 | < 0.05 | < 0.05 | < 0.005 | 0.004 | < 0.002 | < 0.002 | [11 empty] | 16
| 1 | 4 | < 0.001 | < 0.005 | 0.002 | 0.002 | < 0.002 | < 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.005 | < 0.002 | < 0.01 | < 0.05 | < 0.05 | < 0.005 | 0.005 | < 0.002 | < 0.002 | [11 empty] | < 0.005
| 1 | 5 | 0.001 | < 0.005 | 0.002 | 0.002 | 0.002 | 0.002 | 0.002 | 0.003 | < 0.002 | < 0.005 | < 0.002 | < 0.01 | < 0.05 | < 0.05 | < 0.005 | 0.005 | 0.002 | < 0.002 | [11 empty] | 47
| 1 | 6 | 0.001 | < 0.005 | 0.003 | 0.002 | 0.002 | < 0.002 | < 0.002 | 0.003 | < 0.002 | < 0.005 | 0.003 | < 0.01 | < 0.05 | < 0.05 | < 0.005 | 0.003 | < 0.002 | < 0.002
| Mean |  | 1.02: 0.001; 1.04: 0.00225; 1.05: 0.002; 1.06: 0.002; 1.07: 0.002; 1.08: 0.002; 1.09: 0.0028333; 1.12: 0.003; 1.20: 0.004; 1.27: 0.002
| Median |  | 1.02: 0.001; 1.04: 0.002; 1.05: 0.002; 1.06: 0.002; 1.07: 0.002; 1.08: 0.002; 1.09: 0.003; 1.12: 0.003; 1.20: 0.004; 1.27: 0.002
| Std Dev. |  | 1.02: 0; 1.04: 0.0005; 1.05: 0; 1.06: 0; 1.07: 0; 1.08: 0; 1.09: 0.00040825; 1.20: 0.00089443; 1.27: 0
| Rel.Std.Dev. |  | 1.02: 0%; 1.04: 22.22%; 1.05: 0%; 1.06: 0%; 1.07: 0%; 1.08: 0%; 1.09: 14.41%; 1.20: 22.36%; 1.27: 0%
PDM3
| Z-Score (Absolute) |  | 1.01 | 0.67 | 0 | 0 | 0.17 | 0.34 | 0.51 | 1.01 | 0.84 | 0.67 | 0.51 | 3.2 | 23.43 | 23.43 | 0.67 | 2.19 | 0.51 | 0.84
| Analytical results for S in OREAS 627 (Certified Value 10 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 9.98 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.04 | 10.5454 | 10 | 10.35 | > 1 | > 10 | 7.7 | 11.3 | > 10 | 11.4676 | > 10 | 10.1904 | 9.901 | 9.9 | > 5 | > 5 | > 10 | > 5 | 9.25 |  |  |  | 1
| 1 | 2 | 9.81 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.02 | 10.5112 | > 10 | 10.98 | > 1 | > 10 | 7.57 | 11 | > 10 | 11.4643 | > 10 | 10.0831 | 9.948 | 9.8 | > 5 | > 5 | > 10 | > 5 | 8.62 |  |  |  | 26
| 1 | 3 | 9.92 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 7.89 | 10.4292 | 9.8 | 9.93 | > 1 | > 10 | 7.5 | 10.7 | > 10 | 11.7507 | > 10 | 10.2379 | > 10 | 9.9 | > 5 | > 5 | > 10 | > 5 | 8.56 |  |  |  | 16
| 1 | 4 | 10.04 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 7.94 | 10.4135 | > 10 | > 10 | > 1 | > 10 | 7.71 | 11 | > 10 | 11.6712 | > 10 | 10.1687 | 9.836 | 9.9 | > 5 | > 5 | > 10 | > 5 | 8.89 |  |  |  | 10.0035
| 1 | 5 | 10.05 | > 10 | > 10 | 9.82 | > 10 | > 10 | > 10 | 7.95 | 10.3995 | > 10 | 10.58 | > 1 | > 10 | 7.83 | 11.2 | > 10 | 11.1575 | > 10 | 9.8749 | > 10 | 9.8 | > 5 | > 5 | > 10 | > 5 | 9.13 |  |  |  | 48
| 1 | 6 | 10 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.22 | 10.3406 | > 10 | 10.88 | > 1 | > 10 | 7.68 | 11.3 | > 10 | 11.2943 | > 10 | 10.3887 | > 10 | 9.9 | > 5 | > 5 | > 10 | > 5 | 8.9
| Mean |  | 1.01: 9.9667; 1.05: 9.82; 1.09: 8.01; 1.10: 10.4399; 1.11: 9.9; 1.12: 10.544; 1.16: 7.665; 1.17: 11.0833; 1.20: 11.4676; 1.22: 10.1573; 1.23: 9.895; 1.24: 9.8667; 1.30: 8.8917
| Median |  | 1.01: 9.99; 1.05: 9.82; 1.09: 7.985; 1.10: 10.4214; 1.11: 9.9; 1.12: 10.58; 1.16: 7.69; 1.17: 11.1; 1.20: 11.4659; 1.22: 10.1796; 1.23: 9.901; 1.24: 9.9; 1.30: 8.895
| Std Dev. |  | 1.01: 0.089815; 1.09: 0.11662; 1.10: 0.075522; 1.11: 0.14142; 1.12: 0.42395; 1.16: 0.11572; 1.17: 0.23166; 1.20: 0.22262; 1.22: 0.17119; 1.23: 0.056241; 1.24: 0.05164; 1.30: 0.27169
| Rel.Std.Dev. |  | 1.01: 0.9012%; 1.09: 1.456%; 1.10: 0.7234%; 1.11: 1.428%; 1.12: 4.021%; 1.16: 1.51%; 1.17: 2.09%; 1.20: 1.941%; 1.22: 1.685%; 1.23: 0.5684%; 1.24: 0.5234%; 1.30: 3.056%
| PDM3 |  | 1.01: -0.3683%; 1.05: -1.834%; 1.09: -19.93%; 1.10: 4.362%; 1.11: -1.035%; 1.12: 5.403%; 1.16: -23.38%; 1.17: 10.79%; 1.20: 14.64%; 1.22: 1.537%; 1.23: -1.085%; 1.24: -1.368%; 1.30: -11.11%
| Z-Score (Absolute) |  | 0.08 | NA | NA | 0.1 | NA | NA | NA | 2.36 | 0.67 | 0 | 0.8 | NA | NA | 2.79 | 1.48 | NA | 1.96 | NA | 0.32 | 0.01 | 0.04 | NA | NA | NA | NA | 1.26
| Analytical results for Sb in OREAS 627 (Certified Value 142 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 130 | 45.5 | 140.9 | 156.5 | 137.5 | 141.5 | 138.5 | 159 | 148 | 147.54 | 136.7 | 137.17 | 133.3804 | 142.889 | 141.8733 | 118.4 | 142 | 140.9 | 139 | 151.73 | 140 | 143 | 148 | 108 | 133.4 | 151.4 | 146 | 136 |  | 1
| 1 | 2 | 134 | 53.7 | 144.2 | 155 | 143.5 | 139.5 | 141.5 | 158 | 147 | 144.71 | 140.4 | 145.83 | 134.4715 | 138.17 | 141.0553 | 123.1 | 145 | 143.4 | 138 | 154.22 | 147 | 140 | 147 | 116 | 134.9 | 146.8 | 148 | 127 |  | 45
| 1 | 3 | 137 | 48.4 | 131.9 | 155.5 | 140 | 136 | 138 | 158 | 131 | 147 | 131.8 | 143.99 | 131.9374 | 138.728 | 140.8117 | 118.6 | 144 | 142.6 | 137 | 151.4 | 139 | 143 | 149 | 114 | 135.6 | 152.8 | 147 | 135 |  | 16
| 1 | 4 | 133 | 56.4 | 137.4 | 152.5 | 139 | 138 | 137 | 159 | 144 | 144.44 | 135.6 | 145.95 | 133.6339 | 137.191 | 139.8777 | 119.8 | 145 | 143.7 | 140 | 157.51 | 141 | 147 | 150 | 114 | 137.2 | 147.1 | 145 | 140 |  | 142.0611
| 1 | 5 | 128 | 51.8 | 131.7 | 153 | 135 | 137.5 | 135 | 156.5 | 151 | 142.68 | 136.9 | 143.81 | 136.4688 | 139.168 | 139.8409 | 118.7 | 147 | 141.2 | 141 | 154.23 | 147 | 147 | 148 | 113 | 135.5 | 152 | 148 | 144 |  | 49
| 1 | 6 | 137 | 42.5 | 138.8 | 150.5 | 139.5 | 140.5 | 136.5 | 160.5 | 149 | 144.5 | 141.4 | 145.48 | 136.5397 | 136.953 | 142.1584 | 118.3 | 146 | 141.1 | 138 | 154.05 | 145 | 138 | 147 | 116 | 143.4 | 145 | 142 | 136
| Mean |  | 133.1667 | 49.7167 | 137.4833 | 153.8333 | 139.0833 | 138.8333 | 137.75 | 158.5 | 145 | 145.145 | 137.1333 | 143.705 | 134.4053 | 138.8498 | 140.9362 | 119.4833 | 144.8333 | 142.15 | 138.8333 | 153.8567 | 143.1667 | 143 | 148.1667 | 113.5 | 136.6667 | 149.1833 | 146 | 136.3333
| Median |  | 133.5 | 50.1 | 138.1 | 154 | 139.25 | 138.75 | 137.5 | 158.5 | 147.5 | 144.605 | 136.8 | 144.735 | 134.0527 | 138.449 | 140.9335 | 118.65 | 145 | 141.9 | 138.5 | 154.135 | 143 | 143 | 148 | 114 | 135.55 | 149.25 | 146.5 | 136
| Std Dev. |  | 3.656 | 5.2243 | 4.9628 | 2.2286 | 2.8181 | 2.0412 | 2.2079 | 1.3416 | 7.2388 | 1.8086 | 3.4616 | 3.3318 | 1.8196 | 2.1561 | 0.97178 | 1.8519 | 1.7224 | 1.2438 | 1.472 | 2.2006 | 3.6009 | 3.6332 | 1.169 | 2.9496 | 3.5189 | 3.2695 | 2.2804 | 5.6804
| Rel.Std.Dev. |  | 2.745% | 10.51% | 3.61% | 1.449% | 2.026% | 1.47% | 1.603% | 0.8465% | 4.992% | 1.246% | 2.524% | 2.319% | 1.354% | 1.553% | 0.6895% | 1.55% | 1.189% | 0.875% | 1.06% | 1.43% | 2.515% | 2.541% | 0.789% | 2.599% | 2.575% | 2.192% | 1.562% | 4.167%
| PDM3 |  | -6.261% | -65% | -3.222% | 8.287% | -2.096% | -2.272% | -3.035% | 11.57% | 2.069% | 2.171% | -3.469% | 1.157% | -5.389% | -2.261% | -0.7919% | -15.89% | 1.951% | 0.06256% | -2.272% | 8.303% | 0.7782% | 0.6609% | 4.298% | -20.1% | -3.797% | 5.013% | 2.773% | -4.032%
| Z-Score (Absolute) |  | 1.08 | 14.29 | 0.4 | 2.19 | 0.15 | 0.19 | 0.36 | 2.93 | 0.79 | 0.81 | 0.46 | 0.58 | 0.89 | 0.18 | 0.15 | 3.25 | 0.76 | 0.34 | 0.19 | 2.19 | 0.5 | 0.47 | 1.29 | 4.2 | 0.53 | 1.45 | 0.95 | 0.58
| Analytical results for Sc in OREAS 627 (Certified Value 5.13 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  |  | 2
3
| 1 | 1 | 5 | 5 | 5 | 5.5 | 4.4 | 4.5 | 5 | 6.9 | 5 | 5 | 5.3 | 4.2211 | 5.0861 | 4 | 5.8 | 5.21 | 5.6 | 6.6467 | 5 | 5 | 4.8 | 5 | 5.5 | 5 | 5.9 |  |  |  |  | 1
| 1 | 2 | 5 | 5 | 5 | 5.6 | 4.3 | 4.6 | 5.1 | 6.6 | 5.2 | 5 | 5.2 | 4.197 | 5.175 | 4 | 5.9 | 5.21 | 5.9 | 6.6792 | 5 | 5 | 4.9 | 4.8 | 5.6 | 5.1 | 6.4 |  |  |  |  | 46
| 1 | 3 | 5 | 5 | 4.9 | 5.6 | 4.2 | 4.6 | 5.2 | 6.4 | 5.6 | 5 | 5.2 | 4.2101 | 5.0566 | 4 | 5.7 | 5.29 | 5.8 | 6.7755 | 4.9 | 5 | 5.1 | 4.9 | 5.2 | 5.2 | 5.9 |  |  |  |  | 16
| 1 | 4 | 5 | 5 | 5 | 5.5 | 4.4 | 4.3 | 5.1 | 6.5 | 5.5 | 5 | 5.1 | 4.3394 | 4.8754 | 4 | 5.8 | 5.26 | 5.8 | 6.394 | 5.3 | 5 | 4.7 | 5.1 | 5.5 | 5.1 | 5.8 |  |  |  |  | 5.1311
| 1 | 5 | 5 | 5 | 5.1 | 4.8 | 4.4 | 4.2 | 5.1 | 6.9 | 5.2 | 5 | 5.2 | 4.2152 | 4.7848 | 4 | 5.6 | 5.17 | 5.8 | 6.7898 | 5.1 | 6 | 5 | 5.1 | 5.4 | 5 | 6.2 |  |  |  |  | 50
| 1 | 6 | 5 | 5 | 5 | 5.4 | 4.3 | 4.4 | 5.2 | 6.9 | 5.6 | 5 | 5.2 | 4.2759 | 4.7931 | 4 | 5.4 | 5.17 | 6 | 6.3066 | 4.9 | 5 | 5 | 5 | 5.3 | 5.3 | 5.8
| Mean |  | 5 | 5 | 5 | 5.4 | 4.3333 | 4.4333 | 5.1167 | 6.7 | 5.35 | 5 | 5.2 | 4.2431 | 4.9618 | 4 | 5.7 | 5.2183 | 5.8167 | 6.5986 | 5.0333 | 5.1667 | 4.9167 | 4.9833 | 5.4167 | 5.1167 | 6
| Median |  | 5 | 5 | 5 | 5.5 | 4.35 | 4.45 | 5.1 | 6.75 | 5.35 | 5 | 5.2 | 4.2182 | 4.966 | 4 | 5.75 | 5.21 | 5.8 | 6.6629 | 5 | 5 | 4.95 | 5 | 5.45 | 5.1 | 5.9
| Std Dev. |  | 0 | 0 | 0.063246 | 0.30332 | 0.08165 | 0.1633 | 0.075277 | 0.22804 | 0.251 | 0 | 0.063246 | 0.054466 | 0.16562 | 0 | 0.17889 | 0.048339 | 0.13292 | 0.2019 | 0.15055 | 0.40825 | 0.1472 | 0.1169 | 0.1472 | 0.1169 | 0.24495
| Rel.Std.Dev. |  | 0% | 0% | 1.265% | 5.617% | 1.884% | 3.683% | 1.471% | 3.404% | 4.692% | 0% | 1.216% | 1.284% | 3.338% | 0% | 3.138% | 0.9263% | 2.285% | 3.06% | 2.991% | 7.902% | 2.994% | 2.346% | 2.717% | 2.285% | 4.082%
| PDM3 |  | -2.555% | -2.555% | -2.555% | 5.24% | -15.55% | -13.6% | -0.2815% | 30.58% | 4.266% | -2.555% | 1.343% | -17.31% | -3.299% | -22.04% | 11.09% | 1.7% | 13.36% | 28.6% | -1.906% | 0.693% | -4.179% | -2.88% | 5.565% | -0.2815% | 16.93%
| Z-Score (Absolute) |  | NA | NA | 0.44 | 0.67 | 2.3 | 2.02 | 0.12 | 4.3 | 0.53 | NA | 0.12 | 2.55 | 0.55 | NA | 1.51 | 0.17 | 1.84 | 4.02 | 0.35 | NA | 0.67 | 0.49 | 0.72 | 0.12 | 2.35
Results from laboratories 1.01, 1.03, 1.11, 1.17 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Se in OREAS 627 (Indicative Value 2.42 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [6 empty] | 2
2
| 1 | 1 | < 5 | 4.7 | 1 | 9 | 4 | 5 | 7 | 11 | 0.9 | 1 | 2 | 1.5 | 2.2517 | < 2 | 1 | < 10 | 0.72 | 2.1 | 2 | 8 | 7 | 1 | < 10 | [6 empty] | 1
| 1 | 2 | < 5 | 4.6 | 1 | 9 | 5 | 5 | 6 | 11 | 1 | 1.1 | 2 | 1.5 | 2.2233 | < 2 | 1 | < 10 | 0.69 | 2.3 | 2 | 9 | 7 | < 1 | < 10 | [6 empty] | 16
| 1 | 3 | < 5 | 4.8 | < 1 | 8 | 5 | 5 | 7 | 9 | 0.9 | 1 | 1 | 1.5 | 2.3386 | < 2 | 1 | < 10 | 0.67 | 2.2 | 2 | 7 | 7 | < 1 | < 10 | [6 empty] | 16
| 1 | 4 | < 5 | 4.9 | < 1 | 8 | 5 | 5 | 6 | 10 | 1 | 1.1 | 2 | 1.2 | 2.1917 | < 2 | 1 | < 10 | 0.69 | 2 | 2 | 8 | 7 | < 1 | < 10 | [6 empty] | 2.4176
| 1 | 5 | < 5 | 4.4 | 1 | 9 | 5 | 5 | 6 | 10 | 1 | 1 | 2 | 1.4 | 1.9305 | < 2 | 1 | < 10 | 0.71 | 2.3 | 2 | 9 | 7 | < 1 | < 10 | [6 empty] | 9
| 1 | 6 | < 5 | 4.7 | 1 | 8 | 5 | 5 | 6 | 11 | 1 | 1.1 | 2 | 1.4 | 2.2429 | < 2 | 1 | < 10 | 0.71 | 2.3 | 2 | 8 | 8 | < 1 | < 10
| Mean |  |  | 4.6833 | 1 | 8.5 | 4.8333 | 5 | 6.3333 | 10.3333 | 0.96667 | 1.05 | 1.8333 | 1.4167 | 2.1965 |  | 1 |  | 0.69833 | 2.2 | 2 | 8.1667 | 7.1667 | 1
| Median |  |  | 4.7 | 1 | 8.5 | 5 | 5 | 6 | 10.5 | 1 | 1.05 | 2 | 1.45 | 2.2331 |  | 1 |  | 0.7 | 2.25 | 2 | 8 | 7 | 1
| Std Dev. |  |  | 0.17224 | 0 | 0.54772 | 0.40825 | 0 | 0.5164 | 0.8165 | 0.05164 | 0.054772 | 0.40825 | 0.1169 | 0.13921 |  | 0 |  | 0.018348 | 0.12649 | 0 | 0.75277 | 0.40825
| Rel.Std.Dev. |  |  | 3.678% | 0% | 6.444% | 8.447% | 0% | 8.154% | 7.902% | 5.342% | 5.216% | 22.27% | 8.252% | 6.338% |  | 0% |  | 2.627% | 5.75% | 0% | 9.218% | 5.696%
| PDM3 |  |  | 93.71% | -58.64% | 251.6% | 99.92% | 106.8% | 162% | 327.4% | -60.02% | -56.57% | -24.17% | -41.4% | -9.148% |  | -58.64% |  | -71.12% | -9.002% | -17.27% | 237.8% | 196.4% | -58.64%
| Z-Score (Absolute) |  | 0.13 | 1.12 | 0.61 | 2.83 | 1.18 | 1.26 | 1.86 | 3.65 | 0.55 | 0.52 | 0.16 | 0.35 | 0 | 0.54 | 0.54 | 1.26 | 0.67 | 0 | 0.09 | 2.68 | 2.23 | 0.73 | 1.26
| Analytical results for Sm in OREAS 627 (Certified Value 6.24 ppm) | [30 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.22; 1.06: 1.30 | [6 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 5.8; 1.02: 5.92; 1.03: 6.01; 1.04: 4.1117; 1.05: 6.6631; 1.06: 6.22 | [6 empty] | 1
| 1 | 2 | 1.01: 5.9; 1.02: 6.22; 1.03: 6.28; 1.04: 4.1022; 1.05: 6.6327; 1.06: 6.59 | [6 empty] | 8
| 1 | 3 | 1.01: 6.4; 1.02: 6.21; 1.03: 6.09; 1.04: 4.208; 1.05: 6.7581; 1.06: 6.24 | [6 empty] | 16
| 1 | 4 | 1.01: 6.3; 1.02: 6.01; 1.03: 6.31; 1.04: 4.4326; 1.05: 6.5666; 1.06: 6.13 | [6 empty] | 6.2419
| 1 | 5 | 1.01: 6.1; 1.02: 5.81; 1.03: 6.11; 1.04: 3.9479; 1.05: 6.5845; 1.06: 6.51 | [6 empty] | 51
| 1 | 6 | 1.01: 5.9; 1.02: 5.71; 1.03: 6.17; 1.04: 4.23; 1.05: 6.9016; 1.06: 6.21
| Mean |  | 1.01: 6.0667; 1.02: 5.98; 1.03: 6.1617; 1.04: 4.1721; 1.05: 6.6845; 1.06: 6.3167
| Median |  | 1.01: 6; 1.02: 5.965; 1.03: 6.14; 1.04: 4.1599; 1.05: 6.6479; 1.06: 6.23
| Std Dev. |  | 1.01: 0.24221; 1.02: 0.20823; 1.03: 0.11566; 1.04: 0.16204; 1.05: 0.12617; 1.06: 0.1863
| Rel.Std.Dev. |  | 1.01: 3.993%; 1.02: 3.482%; 1.03: 1.877%; 1.04: 3.884%; 1.05: 1.888%; 1.06: 2.949%
| PDM3 |  | 1.01: -2.807%; 1.02: -4.196%; 1.03: -1.285%; 1.04: -33.16%; 1.05: 7.09%; 1.06: 1.198%
| Z-Score (Absolute) |  | 1.01: 0.19; 1.02: 0.54; 1.03: 0.19; 1.04: 7.78; 1.05: 2.28; 1.06: 0.81
| Analytical results for Sn in OREAS 627 (Certified Value 3.42 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | < 10 | 3 | 3.3 | 3.7 | 3.5 | 3.5 | 3.1 | 3.5 | 3.4 | 3.5 | 3.1 | 3.4 | 3.096 | 2.61 | 3.2981 | < 5 | 3.2 | 5.8 | < 10 | 3.5 | 3.5 | 3.5 | 4 | 2.7 | 3.2 | 3.7 | < 10 | 4.1 |  | 1
| 1 | 2 | < 10 | 3 | 3.6 | 3.7 | 3.4 | 3.5 | 3.2 | 3.6 | 3.4 | 3.4 | 3.2 | 3.5 | 3.4291 | 2.78 | 3.3567 | < 5 | 3.1 | 4.7 | < 10 | 3.55 | 3.5 | 3.4 | 4 | 3 | 3.4 | 3.6 | < 10 | 4.7 |  | 12
| 1 | 3 | < 10 | 4 | 3.2 | 3.7 | 3.5 | 3.3 | 3.1 | 3.6 | 3.3 | 3.4 | 3 | 3.5 | 3.5036 | 2.87 | 3.419 | < 5 | 3.3 | 3.4 | < 10 | 3.53 | 3.5 | 3.4 | 4 | 2.9 | 3.2 | 3.7 | < 10 | 3.8 |  | 16
| 1 | 4 | < 10 | 4 | 3.3 | 3.6 | 3.5 | 3.4 | 3.2 | 3.6 | 3.4 | 3.2 | 3.2 | 3.7 | 3.2333 | 5.52 | 3.4442 | < 5 | 3 | 3.9 | < 10 | 3.62 | 3.5 | 3.6 | 4 | 2.7 | 3.3 | 3.8 | < 10 | 3.9 |  | 3.4185
| 1 | 5 | < 10 | 4 | 3.5 | 3.5 | 3 | 3.4 | 3 | 3.4 | 3.5 | 3.2 | 3.2 | 3.5 | 3.3643 | 2.96 | 3.3456 | < 5 | 3.3 | 3.6 | < 10 | 3.55 | 3.7 | 3.5 | 4 | 2.8 | 3.4 | 3.8 | < 10 | 4.2 |  | 52
| 1 | 6 | < 10 | 4 | 3.3 | 3.6 | 3.6 | 3.6 | 3.1 | 3.7 | 3.5 | 3.3 | 3.3 | 3.6 | 3.2816 | 2.92 | 3.2894 | < 5 | 3.1 | 4 | < 10 | 3.49 | 3.6 | 3.5 | 4 | 2.9 | 3.4 | 3.7 | < 10 | 3.8
| Mean |  |  | 3.6667 | 3.3667 | 3.6333 | 3.4167 | 3.45 | 3.1167 | 3.5667 | 3.4167 | 3.3333 | 3.1667 | 3.5333 | 3.318 | 3.2767 | 3.3588 |  | 3.1667 | 4.2333 |  | 3.54 | 3.55 | 3.4833 | 4 | 2.8333 | 3.3167 | 3.7167 |  | 4.0833
| Median |  |  | 4 | 3.3 | 3.65 | 3.5 | 3.45 | 3.1 | 3.6 | 3.4 | 3.35 | 3.2 | 3.5 | 3.3229 | 2.895 | 3.3512 |  | 3.15 | 3.95 |  | 3.54 | 3.5 | 3.5 | 4 | 2.85 | 3.35 | 3.7 |  | 4
| Std Dev. |  |  | 0.5164 | 0.15055 | 0.08165 | 0.2137 | 0.10488 | 0.075277 | 0.10328 | 0.075277 | 0.12111 | 0.10328 | 0.10328 | 0.14609 | 1.106 | 0.062603 |  | 0.12111 | 0.88694 |  | 0.046476 | 0.083666 | 0.075277 | 0 | 0.12111 | 0.098319 | 0.075277 |  | 0.34303
| Rel.Std.Dev. |  |  | 14.08% | 4.472% | 2.247% | 6.255% | 3.04% | 2.415% | 2.896% | 2.203% | 3.633% | 3.261% | 2.923% | 4.403% | 33.75% | 1.864% |  | 3.824% | 20.95% |  | 1.313% | 2.357% | 2.161% | 0% | 4.274% | 2.964% | 2.025% |  | 8.401%
| PDM3 |  |  | 7.259% | -1.517% | 6.284% | -0.05441% | 0.9207% | -8.83% | 4.333% | -0.05441% | -2.492% | -7.368% | 3.358% | -2.941% | -4.15% | -1.746% |  | -7.368% | 23.84% |  | 3.553% | 3.846% | 1.896% | 17.01% | -17.12% | -2.98% | 8.721% |  | 19.45%
| Z-Score (Absolute) |  | 7.29 | NA | 0.31 | 0.93 | 0.08 | 0.08 | 1.47 | 0.62 | 0.08 | 0.47 | 1.24 | 0.47 | 0.54 | 0.73 | 0.35 | 4.34 | 1.24 | 3.72 | 7.29 | 0.5 | 0.54 | 0.23 | NA | 2.79 | 0.54 | 1.32 | 7.29 | 3.02
Results from laboratories 1.02 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Sr in OREAS 627 (Certified Value 91 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 82 | 34.9 | 75 | 97.1 | 95.3 | 84.3 | 84.4 | 89.4 | 105 | 91.07 | 45 | 97.75 | 80.8939 | 59.2 | 96.2289 | 103.3 | 105 | 89 | 89.65 | 87 | 88.9 | 76.4 | 87 | 86.4 | 81.9 | 103.6 | 71.7 | 98 |  | 1
| 1 | 2 | 79 | 36.4 | 78 | 93.3 | 98.7 | 86.1 | 85.7 | 88.5 | 107 | 90.8 | 47 | 104.09 | 80.7106 | 49.62 | 94.7863 | 100.7 | 105 | 96 | 90.25 | 88 | 93.6 | 75.9 | 88 | 87.5 | 83.8 | 103 | 67.1 | 105 |  | 21
| 1 | 3 | 77 | 46 | 74 | 93.4 | 95.6 | 81.4 | 85.3 | 93.6 | 103 | 92.32 | 42 | 100.61 | 84.0599 | 38.84 | 94.7605 | 102.5 | 106 | 95 | 90.81 | 86 | 89 | 80.2 | 81 | 86.3 | 82.6 | 102.9 | 68.1 | 106 |  | 16
| 1 | 4 | 78 | 36.3 | 77 | 92.7 | 99.2 | 88.1 | 85.6 | 86.2 | 104 | 89.27 | 49 | 103.52 | 81.4866 | 55.99 | 94.884 | 101.6 | 107 | 93 | 90.27 | 90 | 87.1 | 79.7 | 86 | 86 | 83.3 | 106.3 | 66.7 | 94 |  | 90.5273
| 1 | 5 | 78 | 36.7 | 75 | 93.2 | 100.5 | 84.5 | 82.7 | 93.1 | 112 | 93.05 | 48 | 103.4 | 81.8898 | 47.36 | 96.3315 | 103.3 | 108 | 96 | 90.16 | 88 | 93.7 | 82.4 | 95 | 85.1 | 85.4 | 102.5 | 66.3 | 104 |  | 53
| 1 | 6 | 83 | 34.9 | 78 | 93.6 | 95.8 | 87 | 84 | 89.2 | 106 | 90.09 | 40 | 104.78 | 81.78 | 40.47 | 95.6291 | 101 | 105 | 98 | 90.18 | 88 | 92.8 | 81.3 | 79 | 88.5 | 85.8 | 100.8 | 67.9 | 102
| Mean |  | 79.5 | 37.5333 | 76.1667 | 93.8833 | 97.5167 | 85.2333 | 84.6167 | 90 | 106.1667 | 91.1 | 45.1667 | 102.3583 | 81.8034 | 48.58 | 95.4367 | 102.0667 | 106 | 94.5 | 90.22 | 87.8333 | 90.85 | 79.3167 | 86 | 86.6333 | 83.8 | 103.1833 | 67.9667 | 101.5
| Median |  | 78.5 | 36.35 | 76 | 93.35 | 97.25 | 85.3 | 84.85 | 89.3 | 105.5 | 90.935 | 46 | 103.46 | 81.6333 | 48.49 | 95.2566 | 102.05 | 105.5 | 95.5 | 90.215 | 88 | 90.9 | 79.95 | 86.5 | 86.35 | 83.55 | 102.95 | 67.5 | 103
| Std Dev. |  | 2.429 | 4.2203 | 1.7224 | 1.6043 | 2.2212 | 2.3746 | 1.1583 | 2.8376 | 3.1885 | 1.3965 | 3.5449 | 2.668 | 1.2018 | 8.1351 | 0.72815 | 1.1361 | 1.2649 | 3.1464 | 0.36922 | 1.3292 | 2.8557 | 2.6286 | 5.6569 | 1.1961 | 1.5401 | 1.7971 | 1.9541 | 4.6368
| Rel.Std.Dev. |  | 3.055% | 11.24% | 2.261% | 1.709% | 2.278% | 2.786% | 1.369% | 3.153% | 3.003% | 1.533% | 7.849% | 2.607% | 1.469% | 16.75% | 0.763% | 1.113% | 1.193% | 3.33% | 0.4092% | 1.513% | 3.143% | 3.314% | 6.578% | 1.381% | 1.838% | 1.742% | 2.875% | 4.568%
| PDM3 |  | -12.18% | -58.54% | -15.86% | 3.707% | 7.721% | -5.848% | -6.529% | -0.5825% | 17.28% | 0.6326% | -50.11% | 13.07% | -9.637% | -46.34% | 5.423% | 12.75% | 17.09% | 4.388% | -0.3395% | -2.976% | 0.3564% | -12.38% | -5.001% | -4.301% | -7.431% | 13.98% | -24.92% | 12.12%
| Z-Score (Absolute) |  | 0.81 | 4.41 | 1.09 | 0.43 | 0.74 | 0.32 | 0.37 | 0.09 | 1.48 | 0.19 | 3.75 | 1.15 | 0.61 | 3.46 | 0.56 | 1.13 | 1.47 | 0.48 | 0.11 | 0.09 | 0.17 | 0.82 | 0.25 | 0.2 | 0.44 | 1.22 | 1.8 | 1.08
| Analytical results for Ta in OREAS 627 (Certified Value 0.79 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 2
2
| 1 | 1 | 8 | 0.3 | 0.9 | 0.84 | 0.82 | 0.88 | 0.72 | 0.79 | 0.75 | 0.78 | 0.6 | 0.77 | 0.74962 | 0.74 | 1.1 | 0.64 | 0.9 | 0.71 | 0.79 | 0.69 | 0.87 | 2.09 | [7 empty] | 1
| 1 | 2 | 8 | 0.3 | 0.9 | 0.84 | 0.84 | 0.88 | 0.72 | 0.83 | 0.73 | 0.79 | 0.6 | 0.81 | 0.72494 | 0.7 | 1 | 0.68 | 0.9 | 0.67 | 0.8 | 2.03 | 0.83 | 1.18 | [7 empty] | 13
| 1 | 3 | 7 | 0.4 | 0.9 | 0.83 | 0.84 | 0.83 | 0.71 | 0.8 | 0.76 | 0.81 | 0.6 | 0.81 | 0.72971 | 0.71 | 0.9 | 0.64 | 1 | 0.65 | 0.83 | 0.58 | 0.82 | 1.33 | [7 empty] | 16
| 1 | 4 | 8 | 0.2 | 0.9 | 0.86 | 0.84 | 0.85 | 0.71 | 0.77 | 0.71 | 0.8 | 0.6 | 0.84 | 0.72941 | 0.71 | 1.1 | 0.65 | 1 | 0.72 | 0.85 | 0.81 | 0.82 | 1.82 | [7 empty] | 0.78976
| 1 | 5 | 8 | 0.3 | 0.9 | 0.82 | 0.67 | 0.82 | 0.71 | 0.78 | 0.74 | 0.83 | 0.6 | 0.8 | 0.73344 | 0.69 | 1.1 | 0.67 | 1 | 0.68 | 0.81 | 0.64 | 0.81 | 2.16 | [7 empty] | 54
| 1 | 6 | 8 | 0.4 | 0.9 | 0.81 | 0.85 | 0.82 | 0.71 | 0.82 | 0.76 | 0.8 | 0.7 | 0.81 | 0.74969 | 0.68 | 1 | 0.65 | 1 | 0.77 | 0.81 | 0.4 | 0.81 | 1.74
| Mean |  | 7.8333 | 0.31667 | 0.9 | 0.83333 | 0.81 | 0.84667 | 0.71333 | 0.79833 | 0.74167 | 0.80167 | 0.61667 | 0.80667 | 0.73613 | 0.705 | 1.0333 | 0.655 | 0.96667 | 0.7 | 0.815 | 0.85833 | 0.82667 | 1.72
| Median |  | 8 | 0.3 | 0.9 | 0.835 | 0.84 | 0.84 | 0.71 | 0.795 | 0.745 | 0.8 | 0.6 | 0.81 | 0.73157 | 0.705 | 1.05 | 0.65 | 1 | 0.695 | 0.81 | 0.665 | 0.82 | 1.78
| Std Dev. |  | 0.40825 | 0.075277 | 0 | 0.017512 | 0.069282 | 0.028048 | 0.005164 | 0.023166 | 0.019408 | 0.017224 | 0.040825 | 0.022509 | 0.010812 | 0.020736 | 0.08165 | 0.016432 | 0.05164 | 0.042895 | 0.021679 | 0.58969 | 0.022509 | 0.39613
| Rel.Std.Dev. |  | 5.212% | 23.77% | 0% | 2.101% | 8.553% | 3.313% | 0.7239% | 2.902% | 2.617% | 2.149% | 6.62% | 2.79% | 1.469% | 2.941% | 7.902% | 2.509% | 5.342% | 6.128% | 2.66% | 68.7% | 2.723% | 23.03%
| PDM3 |  | 891.9% | -59.9% | 13.96% | 5.517% | 2.563% | 7.205% | -9.677% | 1.085% | -6.09% | 1.507% | -21.92% | 2.141% | -6.79% | -10.73% | 30.84% | -17.06% | 22.4% | -11.37% | 3.196% | 8.683% | 4.673% | 117.8%
| Z-Score (Absolute) |  | NA | 4.68 | 0.89 | 0.25 | 0.03 | 0.38 | 0.89 | 0.08 | 0.62 | 0.05 | 1.82 | 0 | 0.67 | 0.97 | 2.17 | 1.45 | 1.53 | 1.02 | 0.08 | 0.49 | 0.19 | 8.73
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Tb in OREAS 627 (Certified Value 0.61 ppm) | [30 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab | [7 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.21; 1.06: 1.22; 1.07: 1.23; 1.08: 1.25; 1.09: 1.26; 1.10: 1.30 | [7 empty] | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*MS | [7 empty] | 2
2
| 1 | 1 | 1.01: 0.7; 1.02: 0.61; 1.03: 0.63; 1.04: 0.48403; 1.05: 0.6; 1.06: 0.76388; 1.07: 0.57; 1.08: 0.61; 1.09: 0.51; 1.10: 0.66 | [7 empty] | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.62; 1.03: 0.66; 1.04: 0.48849; 1.05: 0.7; 1.06: 0.70114; 1.07: 0.56; 1.08: 0.62; 1.09: 0.51; 1.10: 0.69 | [7 empty] | 36
| 1 | 3 | 1.01: 0.7; 1.02: 0.63; 1.03: 0.65; 1.04: 0.48935; 1.05: 0.6; 1.06: 0.76903; 1.07: 0.57; 1.08: 0.68; 1.09: 0.53; 1.10: 0.67 | [7 empty] | 16
| 1 | 4 | 1.01: 0.7; 1.02: 0.61; 1.03: 0.68; 1.04: 0.50781; 1.05: 0.6; 1.06: 0.83884; 1.07: 0.56; 1.08: 0.58; 1.09: 0.54; 1.10: 0.64 | [7 empty] | 0.61451
| 1 | 5 | 1.01: 0.7; 1.02: 0.58; 1.03: 0.66; 1.04: 0.46941; 1.05: 0.6; 1.06: 0.7263; 1.07: 0.59; 1.08: 0.6; 1.09: 0.53; 1.10: 0.68 | [7 empty] | 55
| 1 | 6 | 1.01: 0.7; 1.02: 0.61; 1.03: 0.67; 1.04: 0.48312; 1.05: 0.6; 1.06: 0.82509; 1.07: 0.57; 1.08: 0.67; 1.09: 0.55; 1.10: 0.65
| Mean |  | 1.01: 0.7; 1.02: 0.61; 1.03: 0.65833; 1.04: 0.48703; 1.05: 0.61667; 1.06: 0.77071; 1.07: 0.57; 1.08: 0.62667; 1.09: 0.52833; 1.10: 0.665
| Median |  | 1.01: 0.7; 1.02: 0.61; 1.03: 0.66; 1.04: 0.48626; 1.05: 0.6; 1.06: 0.76646; 1.07: 0.57; 1.08: 0.615; 1.09: 0.53; 1.10: 0.665
| Std Dev. |  | 1.01: 1.2162e-16; 1.02: 0.016733; 1.03: 0.017224; 1.04: 0.012446; 1.05: 0.040825; 1.06: 0.053779; 1.07: 0.010954; 1.08: 0.039833; 1.09: 0.016021; 1.10: 0.018708
| Rel.Std.Dev. |  | 1.01: 1.737e-14%; 1.02: 2.743%; 1.03: 2.616%; 1.04: 2.555%; 1.05: 6.62%; 1.06: 6.978%; 1.07: 1.922%; 1.08: 6.356%; 1.09: 3.032%; 1.10: 2.813%
| PDM3 |  | 1.01: 13.91%; 1.02: -0.7339%; 1.03: 7.131%; 1.04: -20.74%; 1.05: 0.3509%; 1.06: 25.42%; 1.07: -7.243%; 1.08: 1.978%; 1.09: -14.02%; 1.10: 8.216%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.12; 1.03: 0.57; 1.04: 1.86; 1.05: NA; 1.06: 2.16; 1.07: 0.69; 1.08: 0.12; 1.09: 1.28; 1.10: 0.66
Results from laboratories 1.02 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 627 (Certified Value 0.13 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [7 empty] | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [7 empty] | 2
2
| 1 | 1 | 15 | < 0.1 | < 0.5 | 0.14 | 0.12 | 0.1 | 0.14 | 0.16 | 0.14 | < 0.5 | < 0.2 | < 5 | 0.22159 | < 0.1 | 0.1 | < 10 | 0.14 | 0.13 | 0.12 | 0.13 | 0.11 | 0.17 | [7 empty] | 1
| 1 | 2 | 13 | 0.1 | < 0.5 | 0.15 | 0.1 | 0.13 | 0.14 | 0.12 | 0.15 | < 0.5 | < 0.2 | < 5 | 0.19918 | 0.1 | 0.1 | < 10 | 0.12 | 0.11 | 0.12 | 0.15 | 0.14 | 0.13 | [7 empty] | 14
| 1 | 3 | 15 | 0.1 | < 0.5 | 0.15 | 0.09 | 0.11 | 0.14 | 0.11 | 0.15 | < 0.5 | < 0.2 | < 5 | 0.22524 | 0.1 | < 0.1 | < 10 | 0.14 | 0.13 | 0.1 | 0.12 | 0.16 | 0.14 | [7 empty] | 16
| 1 | 4 | 14 | < 0.1 | < 0.5 | 0.12 | 0.1 | 0.11 | 0.14 | 0.08 | 0.13 | < 0.5 | < 0.2 | < 5 | 0.2071 | 0.1 | < 0.1 | < 10 | 0.18 | 0.1 | 0.1 | 0.16 | 0.18 | 0.07 | [7 empty] | 0.12573
| 1 | 5 | 14 | 0.1 | < 0.5 | 0.16 | 0.15 | 0.11 | 0.09 | 0.14 | 0.14 | < 0.5 | < 0.2 | < 5 | 0.19174 | 0.1 | < 0.1 | < 10 | 0.12 | 0.13 | 0.12 | 0.14 | 0.15 | 0.23 | [7 empty] | 56
| 1 | 6 | 14 | 0.1 | < 0.5 | 0.16 | 0.1 | 0.12 | 0.13 | 0.09 | 0.13 | < 0.5 | < 0.2 | < 5 | 0.2105 | < 0.1 | 0.2 | < 10 | 0.16 | 0.13 | 0.13 | 0.14 | 0.17 | 0.09
| Mean |  | 14.1667 | 0.1 |  | 0.14667 | 0.11 | 0.11333 | 0.13 | 0.11667 | 0.14 |  |  |  | 0.20923 | 0.1 | 0.13333 |  | 0.14333 | 0.12167 | 0.115 | 0.14 | 0.15167 | 0.13833
| Median |  | 14 | 0.1 |  | 0.15 | 0.1 | 0.11 | 0.14 | 0.115 | 0.14 |  |  |  | 0.2088 | 0.1 | 0.1 |  | 0.14 | 0.13 | 0.12 | 0.14 | 0.155 | 0.135
| Std Dev. |  | 0.75277 | 0 |  | 0.015055 | 0.021909 | 0.010328 | 0.02 | 0.030111 | 0.0089443 |  |  |  | 0.012823 | 0 | 0.057735 |  | 0.023381 | 0.013292 | 0.012247 | 0.014142 | 0.024833 | 0.057417
| Rel.Std.Dev. |  | 5.314% | 0% |  | 10.27% | 19.92% | 9.113% | 15.38% | 25.81% | 6.389% |  |  |  | 6.129% | 0% | 43.3% |  | 16.31% | 10.92% | 10.65% | 10.1% | 16.37% | 41.51%
| PDM3 |  | 11167% | -20.47% |  | 16.65% | -12.51% | -9.862% | 3.393% | -7.211% | 11.35% |  |  |  | 66.4% | -20.47% | 6.045% |  | 14% | -3.234% | -8.537% | 11.35% | 20.63% | 10.02%
| Z-Score (Absolute) |  | NA | 1.48 | 3.01 | 0.22 | 0.76 | 0.67 | 0.22 | 0.58 | 0.04 | 3.01 | 1.03 | 63.7 | 1.91 | 1.48 | 1.26 | 131.13 | 0.13 | 0.45 | 0.63 | 0.04 | 0.36 | 0
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 627 (Certified Value 9.73 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 3.3 | 7.5 | 9.33 | 11 | 9.04 | 8.63 | 8.25 | 10.1 | 12.23 | 4 | 11.13 | 7.5862 | 10.6572 | 11.2 | 12.3 | 10.43 | 8.8 | 9.6551 | 5.64 | 12 | 7.9 | 9.8 | 11 | 8.84 | [5 empty] | 1
| 1 | 2 | 3.6 | 7.8 | 9.17 | 11.15 | 8.65 | 8.83 | 8.15 | 10.7 | 12.35 | 4.7 | 11.62 | 7.4515 | 10.6877 | 11 | 12.5 | 10.76 | 9.56 | 9.4752 | 5.12 | 11 | 8.2 | 10.1 | 10.8 | 9.07 | [5 empty] | 38
| 1 | 3 | 2.7 | 7.3 | 9.4 | 10.65 | 8.48 | 8.64 | 8.72 | 9.8 | 12.34 | 3.8 | 11.76 | 7.6334 | 10.4072 | 10.7 | 12.5 | 10.55 | 9.03 | 9.7864 | 5.6 | 11 | 9.7 | 9.8 | 10.9 | 8.7 | [5 empty] | 16
| 1 | 4 | 4.2 | 7.7 | 9.37 | 11.3 | 9.06 | 8.06 | 8.04 | 9.3 | 11.79 | 5.1 | 11.64 | 7.8914 | 10.865 | 11.2 | 12.2 | 10.96 | 9.36 | 9.5879 | 6.35 | 10 | 8.3 | 10.3 | 11.3 | 8.68 | [5 empty] | 9.7251
| 1 | 5 | 2 | 7.4 | 9.74 | 10.45 | 8.51 | 8.25 | 8.29 | 11 | 11.6 | 4.4 | 11.71 | 7.1739 | 10.6266 | 11 | 11.7 | 10.83 | 9.56 | 9.8795 | 5.86 | 12 | 7.6 | 10.4 | 11.2 | 12.53 | [5 empty] | 57
| 1 | 6 | 3.3 | 7.2 | 9.58 | 10.7 | 8.41 | 8.26 | 8.08 | 10.9 | 12 | 4 | 12.05 | 7.8127 | 10.3671 | 10.9 | 10.9 | 10.52 | 9.97 | 9.5501 | 7.29 | 11 | 9.2 | 10.3 | 11 | 9.41
| Mean |  | 3.1833 | 7.4833 | 9.4317 | 10.875 | 8.6917 | 8.445 | 8.255 | 10.3 | 12.0517 | 4.3333 | 11.6517 | 7.5915 | 10.6018 | 11 | 12.0167 | 10.675 | 9.38 | 9.6557 | 5.9767 | 11.1667 | 8.4833 | 10.1167 | 11.0333 | 9.5383
| Median |  | 3.3 | 7.45 | 9.385 | 10.85 | 8.58 | 8.445 | 8.2 | 10.4 | 12.115 | 4.2 | 11.675 | 7.6098 | 10.6419 | 11 | 12.25 | 10.655 | 9.46 | 9.6215 | 5.75 | 11 | 8.25 | 10.2 | 11 | 8.955
| Std Dev. |  | 0.75741 | 0.23166 | 0.20034 | 0.32673 | 0.2884 | 0.29697 | 0.24712 | 0.67823 | 0.30954 | 0.49666 | 0.29903 | 0.25871 | 0.18619 | 0.18974 | 0.62102 | 0.20599 | 0.41766 | 0.152 | 0.75722 | 0.75277 | 0.80353 | 0.26394 | 0.18619 | 1.4908
| Rel.Std.Dev. |  | 23.79% | 3.096% | 2.124% | 3.004% | 3.318% | 3.516% | 2.994% | 6.585% | 2.568% | 11.46% | 2.566% | 3.408% | 1.756% | 1.725% | 5.168% | 1.93% | 4.453% | 1.574% | 12.67% | 6.741% | 9.472% | 2.609% | 1.688% | 15.63%
| PDM3 |  | -67.27% | -23.05% | -3.017% | 11.82% | -10.63% | -13.16% | -15.12% | 5.911% | 23.92% | -55.44% | 19.81% | -21.94% | 9.015% | 13.11% | 23.56% | 9.767% | -3.549% | -0.714% | -38.54% | 14.82% | -12.77% | 4.026% | 13.45% | -1.921%
| Z-Score (Absolute) |  | 3.3 | 1.09 | 0.09 | 0.66 | 0.47 | 0.59 | 0.69 | 0.36 | 1.26 | 2.71 | 1.06 | 1.03 | 0.52 | 0.72 | 1.25 | 0.55 | 0.11 | 0.03 | 1.86 | 0.81 | 0.57 | 0.27 | 0.74 | 0.03
| Analytical results for Ti in OREAS 627 (Certified Value 0.137 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.13 | 0.136 | 0.148 | 0.148 | 0.148 | 0.142 | 0.15 | 0.14 | 0.1384 | 0.127 | 0.1333 | 0.14409 | 0.1 | 0.14025 | 0.14 | 0.132 | 0.137 | 0.1265 | 0.14 | 0.1355 | 0.125 | 0.13 | 0.13 | 0.12 | 0.15 | 0.129 | 0.13 |  |  | 1
| 1 | 2 | 0.12 | 0.14 | 0.15 | 0.148 | 0.144 | 0.142 | 0.156 | 0.14 | 0.1404 | 0.128 | 0.1387 | 0.14435 | 0.103 | 0.13299 | 0.14 | 0.132 | 0.139 | 0.1273 | 0.14 | 0.12889 | 0.127 | 0.13 | 0.13 | 0.11 | 0.16 | 0.129 | 0.14 |  |  | 39
| 1 | 3 | 0.12 | 0.138 | 0.152 | 0.145 | 0.145 | 0.139 | 0.154 | 0.13 | 0.1409 | 0.122 | 0.1352 | 0.1424 | 0.1 | 0.13503 | 0.16 | 0.13 | 0.138 | 0.1266 | 0.13 | 0.13309 | 0.128 | 0.14 | 0.13 | 0.12 | 0.14 | 0.13 | 0.14 |  |  | 16
| 1 | 4 | 0.13 | 0.139 | 0.148 | 0.145 | 0.149 | 0.136 | 0.153 | 0.14 | 0.1394 | 0.132 | 0.1386 | 0.14457 | < 0.1 | 0.13884 | 0.14 | 0.13 | 0.14 | 0.1258 | 0.14 | 0.12836 | 0.131 | 0.13 | 0.13 | 0.11 | 0.14 | 0.132 | 0.13 |  |  | 0.13659
| 1 | 5 | 0.13 | 0.142 | 0.149 | 0.133 | 0.148 | 0.139 | 0.154 | 0.14 | 0.1398 | 0.13 | 0.1343 | 0.14726 | < 0.1 | 0.13563 | 0.16 | 0.132 | 0.142 | 0.1238 | 0.14 | 0.12767 | 0.129 | 0.13 | 0.13 | 0.12 | 0.14 | 0.13 | 0.13 |  |  | 58
| 1 | 6 | 0.13 | 0.135 | 0.148 | 0.144 | 0.144 | 0.14 | 0.15 | 0.15 | 0.1369 | 0.135 | 0.1392 | 0.14558 | < 0.1 | 0.13639 | 0.14 | 0.13 | 0.14 | 0.1233 | 0.14 | 0.1324 | 0.129 | 0.13 | 0.13 | 0.12 | 0.14 | 0.128 | 0.13
| Mean |  | 0.12667 | 0.13833 | 0.14917 | 0.14383 | 0.14633 | 0.13967 | 0.15283 | 0.14 | 0.1393 | 0.129 | 0.13655 | 0.14471 | 0.101 | 0.13652 | 0.14667 | 0.131 | 0.13933 | 0.12555 | 0.13833 | 0.13098 | 0.12817 | 0.13167 | 0.13 | 0.11667 | 0.145 | 0.12967 | 0.13333
| Median |  | 0.13 | 0.1385 | 0.1485 | 0.145 | 0.1465 | 0.1395 | 0.1535 | 0.14 | 0.1396 | 0.129 | 0.1369 | 0.14446 | 0.1 | 0.13601 | 0.14 | 0.131 | 0.1395 | 0.12615 | 0.14 | 0.13065 | 0.1285 | 0.13 | 0.13 | 0.12 | 0.14 | 0.1295 | 0.13
| Std Dev. |  | 0.005164 | 0.002582 | 0.0016021 | 0.0055648 | 0.0022509 | 0.0022509 | 0.0024014 | 0.0063246 | 0.001456 | 0.0044721 | 0.0025805 | 0.0016219 | 0.0017321 | 0.0026389 | 0.010328 | 0.0010954 | 0.0017512 | 0.0016282 | 0.0040825 | 0.0031333 | 0.0020412 | 0.0040825 | 0 | 0.005164 | 0.0083666 | 0.0013663 | 0.005164
| Rel.Std.Dev. |  | 4.077% | 1.866% | 1.074% | 3.869% | 1.538% | 1.612% | 1.571% | 4.518% | 1.045% | 3.467% | 1.89% | 1.121% | 1.715% | 1.933% | 7.042% | 0.8362% | 1.257% | 1.297% | 2.951% | 2.392% | 1.593% | 3.101% | 0% | 4.426% | 5.77% | 1.054% | 3.873%
| PDM3 |  | -7.268% | 1.273% | 9.204% | 5.3% | 7.13% | 2.249% | 11.89% | 2.493% | 1.981% | -5.56% | -0.03231% | 5.94% | -26.06% | -0.05343% | 7.374% | -4.095% | 2.005% | -8.085% | 1.273% | -4.106% | -6.17% | -3.607% | -4.828% | -14.59% | 6.154% | -5.072% | -2.387%
| Z-Score (Absolute) |  | 0.97 | 0.17 | 1.24 | 0.71 | 0.96 | 0.31 | 1.6 | 0.34 | 0.27 | 0.74 | 0 | 0.8 | 5.98 | 0 | 0.99 | 0.54 | 0.27 | 1.08 | 0.17 | 0.55 | 0.82 | 0.48 | 0.64 | 1.95 | 0.83 | 0.67 | 0.32
| Analytical results for Tl in OREAS 627 (Certified Value 14.2 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | [6 empty] | 1
2
| 1 | 1 | 14 | 14 | 14.3 | 15.05 | 15.1 | 13.3 | 15 | 14.5 | 14.42 | 15.3 | 14.66 | 12.97 | 13.372 | 14.4477 | 14.7 | 13.5 | 14.24 | 13.64 | 14.23 | 13.6 | 12.95 | 13.39 | 15.88 | [6 empty] | 1
| 1 | 2 | 14.1 | 14.3 | 14.25 | 15.4 | 14.7 | 13.65 | 14.95 | 14.9 | 14.45 | 15.1 | 15.18 | 12.5077 | 15.049 | 14.2987 | 14.8 | 13.8 | 14.396 | 14.81 | 13.63 | 14.5 | 12.44 | 13.11 | 15.22 | [6 empty] | 0
| 1 | 3 | 14.1 | 13.8 | 14.35 | 15.65 | 14.1 | 12.9 | 15.25 | 15.6 | 14.13 | 14.3 | 14.99 | 12.7016 | 14.076 | 14.4059 | 15.1 | 13.6 | 14.324 | 13.37 | 13.48 | 14.6 | 12.89 | 13.27 | 14.91 | [6 empty] | 16
| 1 | 4 | 14.5 | 14 | 13.95 | 15.2 | 15.1 | 12.9 | 15.3 | 14.2 | 14.1 | 15 | 15.46 | 12.5629 | 13.827 | 14.3754 | 14.8 | 13.4 | 14.877 | 14.02 | 13.24 | 14.2 | 13.25 | 13.11 | 15.44 | [6 empty] | 14.2142
| 1 | 5 | 14 | 13.8 | 14.35 | 13.2 | 13.75 | 12.9 | 14.95 | 15.9 | 14.18 | 14.6 | 15.07 | 12.5786 | 13.673 | 14.4814 | 14.6 | 13.6 | 14.68 | 14.36 | 14.09 | 13.9 | 12.91 | 13.22 | 15.87 | [6 empty] | 59
| 1 | 6 | 14.1 | 14.3 | 14.1 | 15 | 15.05 | 12.75 | 15.15 | 14.7 | 14.16 | 15.4 | 15.58 | 12.764 | 14.007 | 14.2734 | 14.9 | 11.1 | 14.252 | 14.24 | 14.34 | 14.4 | 13.53 | 12.73 | 14.69
| Mean |  | 14.1333 | 14.0333 | 14.2167 | 14.9167 | 14.6333 | 13.0667 | 15.1 | 14.9667 | 14.24 | 14.95 | 15.1567 | 12.6808 | 14.0007 | 14.3804 | 14.8167 | 13.1667 | 14.4615 | 14.0733 | 13.835 | 14.2 | 12.995 | 13.1383 | 15.335
| Median |  | 14.1 | 14 | 14.275 | 15.125 | 14.875 | 12.9 | 15.075 | 14.8 | 14.17 | 15.05 | 15.125 | 12.6401 | 13.917 | 14.3907 | 14.8 | 13.55 | 14.36 | 14.13 | 13.86 | 14.3 | 12.93 | 13.165 | 15.33
| Std Dev. |  | 0.18619 | 0.22509 | 0.16021 | 0.87445 | 0.57937 | 0.3401 | 0.15492 | 0.65625 | 0.15375 | 0.42308 | 0.33279 | 0.17046 | 0.57227 | 0.081867 | 0.17224 | 1.0211 | 0.25943 | 0.51729 | 0.4468 | 0.38471 | 0.36866 | 0.22613 | 0.49074
| Rel.Std.Dev. |  | 1.317% | 1.604% | 1.127% | 5.862% | 3.959% | 2.603% | 1.026% | 4.385% | 1.08% | 2.83% | 2.196% | 1.344% | 4.087% | 0.5693% | 1.162% | 7.755% | 1.794% | 3.676% | 3.229% | 2.709% | 2.837% | 1.721% | 3.2%
| PDM3 |  | -0.5691% | -1.273% | 0.0172% | 4.942% | 2.949% | -8.073% | 6.232% | 5.294% | 0.1814% | 5.176% | 6.63% | -10.79% | -1.502% | 1.169% | 4.238% | -7.37% | 1.74% | -0.9912% | -2.668% | -0.1001% | -8.577% | -7.569% | 7.885%
| Z-Score (Absolute) |  | 0.09 | 0.21 | 0 | 0.79 | 0.47 | 1.29 | 0.99 | 0.84 | 0.03 | 0.82 | 1.06 | 1.73 | 0.24 | 0.18 | 0.67 | 1.18 | 0.28 | 0.16 | 0.43 | 0.02 | 1.37 | 1.21 | 1.26
| Analytical results for Tm in OREAS 627 (Certified Value 0.21 ppm) | [30 empty] | Certified
Tm | Round | Replicate | 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | [6 empty] | 1
| No | No | 1.02: 1.02; 1.03: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.22; 1.07: 1.30 | [6 empty] | ppm
|  |  | 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*MS; 1.06: 4A*MS; 1.07: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 1.02: 0.2; 1.03: 0.21; 1.04: 0.2; 1.05: 0.15513; 1.06: 0.20329; 1.07: 0.21 | [6 empty] | 1
| 1 | 2 | 1.02: 0.2; 1.03: 0.2; 1.04: 0.2; 1.05: 0.15143; 1.06: 0.20765; 1.07: 0.23 | [6 empty] | 9
| 1 | 3 | 1.02: 0.2; 1.03: 0.21; 1.04: 0.2; 1.05: 0.15605; 1.06: 0.21439; 1.07: 0.21 | [6 empty] | 16
| 1 | 4 | 1.02: 0.2; 1.03: 0.2; 1.04: 0.2; 1.05: 0.15807; 1.06: 0.2309; 1.07: 0.21 | [6 empty] | 0.20606
| 1 | 5 | 1.02: 0.2; 1.03: 0.2; 1.04: 0.19; 1.05: 0.15306; 1.06: 0.20356; 1.07: 0.22 | [6 empty] | 60
| 1 | 6 | 1.02: 0.2; 1.03: 0.2; 1.04: 0.21; 1.05: 0.15433; 1.06: 0.25886; 1.07: 0.21
| Mean |  | 1.02: 0.2; 1.03: 0.20333; 1.04: 0.2; 1.05: 0.15468; 1.06: 0.21977; 1.07: 0.215
| Median |  | 1.02: 0.2; 1.03: 0.2; 1.04: 0.2; 1.05: 0.15473; 1.06: 0.21102; 1.07: 0.21
| Std Dev. |  | 1.02: 3.0405e-17; 1.03: 0.005164; 1.04: 0.0063246; 1.05: 0.0023171; 1.06: 0.021734; 1.07: 0.0083666
| Rel.Std.Dev. |  | 1.02: 1.52e-14%; 1.03: 2.54%; 1.04: 3.162%; 1.05: 1.498%; 1.06: 9.889%; 1.07: 3.891%
| PDM3 |  | 1.02: -2.94%; 1.03: -1.322%; 1.04: -2.94%; 1.05: -24.94%; 1.06: 6.656%; 1.07: 4.34%
| Z-Score (Absolute) |  | 1.02: 0.15; 1.03: 0.15; 1.04: 0.15; 1.05: 4.22; 1.06: 1.63; 1.07: 1.2
| Analytical results for U in OREAS 627 (Certified Value 4.83 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 5 | 4.5 | 4.9 | 5 | 4.6 | 4.5 | 4.3 | 4.98 | 5.11 | 4.9 | 4.76 | 4.0369 | 4.9851 | 5.02 | 4.7 | 5.15 | 4.61 | 4.1007 | 4.84 | < 20 | 4.82 | 4.6 | 4.9 | < 10 | [5 empty] | 1
| 1 | 2 | 5.1 | 4.6 | 4.9 | 5.2 | 5.2 | 4.7 | 4.2 | 5.1 | 5.15 | 4.9 | 5.11 | 3.9973 | 5.0139 | 5.13 | 4.7 | 5.25 | 4.99 | 4.379 | 4.87 | < 20 | 4.9 | 4.2 | 5 | < 10 | [5 empty] | 41
| 1 | 3 | 5.1 | 4.2 | 5 | 5 | 4.7 | 4.5 | 4.6 | 4.86 | 5.2 | 4.8 | 5.11 | 4.2108 | 4.9757 | 5.05 | 4.7 | 5.21 | 4.68 | 4.1007 | 4.56 | < 20 | 5.11 | 4.7 | 5 | < 10 | [5 empty] | 16
| 1 | 4 | 5.2 | 4.6 | 4.8 | 5 | 4.5 | 4.4 | 4.3 | 4.78 | 5.02 | 4.8 | 5.24 | 4.0892 | 4.9732 | 5.15 | 4.7 | 5.35 | 4.71 | 4.4441 | 4.82 | < 20 | 4.89 | 4.7 | 4.9 | < 10 | [5 empty] | 4.8287
| 1 | 5 | 5.1 | 4.5 | 4.8 | 4.7 | 4.3 | 4.6 | 4.2 | 5.16 | 5.43 | 4.7 | 5.06 | 3.9161 | 5.0247 | 5.16 | 4.6 | 5.35 | 4.82 | 4.1682 | 4.82 | < 20 | 4.87 | 4.6 | 4.9 | < 10 | [5 empty] | 61
| 1 | 6 | 5.1 | 4.5 | 4.8 | 4.9 | 4.8 | 4.4 | 4.3 | 4.93 | 5.07 | 4.9 | 5.18 | 4.0191 | 4.9594 | 5.01 | 4.6 | 5.14 | 4.88 | 4.3559 | 4.91 | < 20 | 5.09 | 4.8 | 4.8 | < 10
| Mean |  | 5.1 | 4.4833 | 4.8667 | 4.9667 | 4.6833 | 4.5167 | 4.3167 | 4.9683 | 5.1633 | 4.8333 | 5.0767 | 4.0449 | 4.9887 | 5.0867 | 4.6667 | 5.2417 | 4.7817 | 4.2581 | 4.8033 |  | 4.9467 | 4.6 | 4.9167
| Median |  | 5.1 | 4.5 | 4.85 | 5 | 4.65 | 4.5 | 4.3 | 4.955 | 5.13 | 4.85 | 5.11 | 4.028 | 4.9804 | 5.09 | 4.7 | 5.23 | 4.765 | 4.262 | 4.83 |  | 4.895 | 4.65 | 4.9
| Std Dev. |  | 0.063246 | 0.1472 | 0.08165 | 0.1633 | 0.30605 | 0.1169 | 0.1472 | 0.14345 | 0.14473 | 0.08165 | 0.16741 | 0.099041 | 0.025338 | 0.067725 | 0.05164 | 0.093041 | 0.14106 | 0.15256 | 0.12404 |  | 0.12209 | 0.20976 | 0.075277
| Rel.Std.Dev. |  | 1.24% | 3.283% | 1.678% | 3.288% | 6.535% | 2.588% | 3.41% | 2.887% | 2.803% | 1.689% | 3.298% | 2.449% | 0.5079% | 1.331% | 1.107% | 1.775% | 2.95% | 3.583% | 2.582% |  | 2.468% | 4.56% | 1.531%
| PDM3 |  | 5.617% | -7.153% | 0.7852% | 2.856% | -3.011% | -6.463% | -10.6% | 2.891% | 6.929% | 0.09492% | 5.134% | -16.23% | 3.312% | 5.341% | -3.357% | 8.551% | -0.9751% | -11.82% | -0.5264% |  | 2.442% | -4.737% | 1.821%
| Z-Score (Absolute) |  | 0.7 | 1.37 | 0.08 | 0.25 | 0.7 | 1.25 | 1.92 | 0.26 | 0.91 | 0.2 | 0.62 | 2.83 | 0.32 | 0.65 | 0.75 | 1.17 | 0.37 | 2.12 | 0.3 | 17.08 | 0.18 | 0.98 | 0.08 | 0.36
| Analytical results for V in OREAS 627 (Certified Value 23.8 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 1
2
| 1 | 1 | 25 | 21 | 23 | 24 | 22 | 23 | 23 | 23 | 24.8 | 23 | 25 | 25 | 23.0082 | 24.269 | 24.0464 | 37.04 | 24 | 23 | 25.25 | 24 | 23 | 26.77 | 23 | 25 | 24 | 25 | 24 | 26 | 22 | 1
| 1 | 2 | 25 | 21 | 24 | 24 | 22 | 22 | 23 | 24 | 25.2 | 23 | 25 | 24 | 23.6526 | 24.305 | 24.1662 | 36.57 | 24 | 23 | 25.15 | 24 | 25 | 25.8 | 23 | 25 | 25 | 23 | 24 | 26 | 22 | 23
| 1 | 3 | 25 | 22 | 24 | 24 | 22 | 22 | 22 | 23 | 24.3 | 23 | 24 | 23 | 23.5701 | 24.689 | 23.3366 | 36.61 | 25 | 23 | 24.67 | 24 | 24 | 26.18 | 23 | 25 | 25 | 24 | 23 | 26 | 23 | 16
| 1 | 4 | 26 | 21 | 24 | 23 | 21 | 23 | 22 | 24 | 24.5 | 22 | 25 | 23 | 23.5325 | 23.752 | 23.3902 | 37.4 | 24 | 23 | 24.86 | 23 | 24 | 25.22 | 23 | 25 | 24 | 25 | 23 | 26 | 21 | 23.7633
| 1 | 5 | 26 | 21 | 24 | 23 | 22 | 23 | 22 | 23 | 26.2 | 22 | 25 | 23 | 24.3548 | 24.235 | 23.1221 | 37.71 | 24 | 24 | 24.98 | 23 | 25 | 25.64 | 24 | 25 | 23 | 25 | 23 | 26 | 21 | 62
| 1 | 6 | 26 | 21 | 23 | 24 | 21 | 23 | 22 | 23 | 26.1 | 22 | 26 | 23 | 23.2275 | 24.619 | 23.3485 | 38.81 | 24 | 21 | 24.86 | 23 | 25 | 26.13 | 23 | 25 | 25 | 25 | 24 | 25 | 21
| Mean |  | 25.5 | 21.1667 | 23.6667 | 23.6667 | 21.6667 | 22.6667 | 22.3333 | 23.3333 | 25.1833 | 22.5 | 25 | 23.5 | 23.5576 | 24.3115 | 23.5683 | 37.3567 | 24.1667 | 22.8333 | 24.9617 | 23.5 | 24.3333 | 25.9567 | 23.1667 | 25 | 24.3333 | 24.5 | 23.5 | 25.8333 | 21.6667
| Median |  | 25.5 | 21 | 24 | 24 | 22 | 23 | 22 | 23 | 25 | 22.5 | 25 | 23 | 23.5513 | 24.287 | 23.3693 | 37.22 | 24 | 23 | 24.92 | 23.5 | 24.5 | 25.965 | 23 | 25 | 24.5 | 25 | 23.5 | 26 | 21.5
| Std Dev. |  | 0.54772 | 0.40825 | 0.5164 | 0.5164 | 0.5164 | 0.5164 | 0.5164 | 0.5164 | 0.8085 | 0.54772 | 0.63246 | 0.83666 | 0.45961 | 0.33401 | 0.42869 | 0.83851 | 0.40825 | 0.98319 | 0.21198 | 0.54772 | 0.8165 | 0.53084 | 0.40825 | 0 | 0.8165 | 0.83666 | 0.54772 | 0.40825 | 0.8165
| Rel.Std.Dev. |  | 2.148% | 1.929% | 2.182% | 2.182% | 2.383% | 2.278% | 2.312% | 2.213% | 3.21% | 2.434% | 2.53% | 3.56% | 1.951% | 1.374% | 1.819% | 2.245% | 1.689% | 4.306% | 0.8492% | 2.331% | 3.355% | 2.045% | 1.762% | 0% | 3.355% | 3.415% | 2.331% | 1.58% | 3.768%
| PDM3 |  | 7.308% | -10.93% | -0.4067% | -0.4067% | -8.823% | -4.615% | -6.018% | -1.809% | 5.976% | -5.316% | 5.204% | -1.108% | -0.8656% | 2.307% | -0.8205% | 57.2% | 1.697% | -3.913% | 5.043% | -1.108% | 2.399% | 9.23% | -2.511% | 5.204% | 2.399% | 3.1% | -1.108% | 8.711% | -8.823%
| Z-Score (Absolute) |  | 1.48 | 2.02 | 0 | 0 | 1.62 | 0.81 | 1.08 | 0.27 | 1.23 | 0.94 | 1.08 | 0.13 | 0.09 | 0.52 | 0.08 | 11.08 | 0.4 | 0.67 | 1.05 | 0.13 | 0.54 | 1.85 | 0.4 | 1.08 | 0.54 | 0.67 | 0.13 | 1.75 | 1.62
| Analytical results for W in OREAS 627 (Certified Value 8.25 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | < 10 | 7.9 | 7.9 | 9.4 | 8.7 | 9 | 8.6 | 8.4 | 8.5 | 8.6 | 7.7 | 8.4 | < 10 | 7.6882 | < 5 | 8.1 | 7.6 | 8.66 | 8.3 | 8 | 7 | 7 | 7.1 | 8.7 | 378 | < 20 |  |  |  | 1
| 1 | 2 | < 10 | 8 | 7.9 | 9.6 | 9.3 | 9 | 8.7 | 8.8 | 7.8 | 8.7 | 7.6 | 8.7 | < 10 | 7.6709 | < 5 | 7.8 | 7.6 | 8.07 | 8.9 | 7.6 | 7 | 7 | 7 | 8.5 | 383 | < 20 |  |  |  | 43
| 1 | 3 | < 10 | 8.1 | 7.5 | 9.3 | 9 | 8.5 | 8.6 | 8.6 | 8.3 | 9.1 | 7.3 | 8.6 | < 10 | 7.6393 | < 5 | 8 | 7.6 | 7.93 | 8.3 | 7.8 | 7 | 7 | 7.3 | 8.5 | 375 | < 20 |  |  |  | 16
| 1 | 4 | < 10 | 6 | 7.9 | 9.6 | 9 | 8.8 | 8.2 | 8.7 | 8.5 | 8.7 | 7.6 | 8.8 | < 10 | 7.5218 | < 5 | 8 | 7.7 | 8.31 | 8.3 | 8.1 | 7 | 7 | 7.1 | 9 | 382 | < 20 |  |  |  | 8.2543
| 1 | 5 | < 10 | 7.9 | 7.9 | 9.2 | 7.8 | 8.4 | 8.3 | 8.5 | 8.6 | 8.7 | 7.4 | 8.4 | < 10 | 7.7489 | < 5 | 7.8 | 7.5 | 8.26 | 8.7 | 7.9 | 8 | 7 | 7.2 | 8.8 | 386 | < 20 |  |  |  | 63
| 1 | 6 | < 10 | 8.5 | 8 | 9 | 8.8 | 8.5 | 8.3 | 8.6 | 7.6 | 8.7 | 7.9 | 8.9 | < 10 | 7.4836 | < 5 | 7.9 | 6.1 | 8.05 | 8.4 | 7.9 | 7 | 7 | 7.3 | 8.6 | 374 | < 20
| Mean |  |  | 7.7333 | 7.85 | 9.35 | 8.7667 | 8.7 | 8.45 | 8.6 | 8.2167 | 8.75 | 7.5833 | 8.6333 |  | 7.6255 |  | 7.9333 | 7.35 | 8.2133 | 8.4833 | 7.8833 | 7.1667 | 7 | 7.1667 | 8.6833 | 379.6667
| Median |  |  | 7.95 | 7.9 | 9.35 | 8.9 | 8.65 | 8.45 | 8.6 | 8.4 | 8.7 | 7.6 | 8.65 |  | 7.6551 |  | 7.95 | 7.6 | 8.165 | 8.35 | 7.9 | 7 | 7 | 7.15 | 8.65 | 380
| Std Dev. |  |  | 0.87788 | 0.17607 | 0.23452 | 0.5164 | 0.26833 | 0.20736 | 0.14142 | 0.41673 | 0.17607 | 0.2137 | 0.20656 |  | 0.10226 |  | 0.12111 | 0.61563 | 0.26021 | 0.25626 | 0.17224 | 0.40825 | 0 | 0.12111 | 0.19408 | 4.761
| Rel.Std.Dev. |  |  | 11.35% | 2.243% | 2.508% | 5.89% | 3.084% | 2.454% | 1.644% | 5.072% | 2.012% | 2.818% | 2.393% |  | 1.341% |  | 1.527% | 8.376% | 3.168% | 3.021% | 2.185% | 5.696% | 0% | 1.69% | 2.235% | 1.254%
| PDM3 |  |  | -6.312% | -4.898% | 13.27% | 6.207% | 5.399% | 2.371% | 4.188% | -0.4562% | 6.005% | -8.129% | 4.592% |  | -7.619% |  | -3.889% | -10.96% | -0.4966% | 2.774% | -4.494% | -13.18% | -15.2% | -13.18% | 5.197% | 4500%
| Z-Score (Absolute) |  | 4.25 | 0.64 | 0.48 | 1.5 | 0.73 | 0.64 | 0.31 | 0.51 | 0 | 0.71 | 0.84 | 0.55 | 4.25 | 0.78 | 7.56 | 0.37 | 1.14 | 0 | 0.35 | 0.44 | NA | NA | 1.39 | 0.62 | 491.12 | 2.36
Results from laboratories 1.24 and 1.25 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 627 (Certified Value 13.9 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  |  | 1
1
| 1 | 1 | 13 | 15.7 | 12.8 | 13 | 23.3 | 11.2 | 12.1 | 12.3 | 15.5 | 14.29 | 13.6 | 14.09 | 11.1659 | 14.8074 | 15.2 | 16.2 | 11.24 | 13.3 | 13.4 | 13.1 | 16 | 12.4 | 12.5 | 15.7 | 17.4 | 15 |  |  |  | 1
| 1 | 2 | 13 | 16 | 13.4 | 13 | 16.7 | 11.3 | 12.6 | 12.9 | 15.3 | 14.07 | 14 | 14.62 | 11.0031 | 14.5396 | 15.2 | 16.4 | 11.25 | 13.22 | 14.4 | 12.9 | 16 | 12.8 | 11.7 | 15.2 | 17.5 | 16.3 |  |  |  | 44
| 1 | 3 | 13 | 16.2 | 11.7 | 13 | 16.4 | 10.8 | 13 | 12.9 | 14.9 | 14.51 | 13.6 | 14.35 | 10.966 | 14.2918 | 16.2 | 16.4 | 11.15 | 13.04 | 14.1 | 12.5 | 16 | 14.1 | 12.6 | 16.2 | 17.5 | 15.3 |  |  |  | 16
| 1 | 4 | 14 | 14.7 | 12.8 | 13.4 | 16.4 | 11.4 | 11.9 | 12.2 | 15.2 | 14.07 | 13.5 | 14.52 | 11.5628 | 14.9041 | 15.4 | 15.3 | 11.08 | 13.43 | 13.1 | 12.8 | 15 | 11.3 | 12.5 | 16.1 | 17.5 | 14.6 |  |  |  | 13.8965
| 1 | 5 | 14 | 16.1 | 12 | 13.7 | 14.3 | 11.6 | 11.5 | 11.9 | 16 | 13.65 | 13.7 | 14.48 | 10.8953 | 14.341 | 15.7 | 15 | 11.09 | 13.29 | 14.5 | 13 | 17 | 12.5 | 12.5 | 16.1 | 17.1 | 15.7 |  |  |  | 64
| 1 | 6 | 13 | 16.1 | 12.1 | 13.6 | 16.3 | 10.8 | 12.2 | 13 | 15.9 | 13.88 | 13.9 | 14.52 | 11.3411 | 14.551 | 15.5 | 13.9 | 11.15 | 13.29 | 14.5 | 13.2 | 15 | 13.8 | 13.3 | 15.9 | 17.7 | 15
| Mean |  | 13.3333 | 15.8 | 12.4667 | 13.2833 | 17.2333 | 11.1833 | 12.2167 | 12.5333 | 15.4667 | 14.0783 | 13.7167 | 14.43 | 11.1557 | 14.5725 | 15.5333 | 15.5333 | 11.16 | 13.2617 | 14 | 12.9167 | 15.8333 | 12.8167 | 12.5167 | 15.8667 | 17.45 | 15.3167
| Median |  | 13 | 16.05 | 12.45 | 13.2 | 16.4 | 11.25 | 12.15 | 12.6 | 15.4 | 14.07 | 13.65 | 14.5 | 11.0845 | 14.5453 | 15.45 | 15.75 | 11.15 | 13.29 | 14.25 | 12.95 | 16 | 12.65 | 12.5 | 16 | 17.5 | 15.15
| Std Dev. |  | 0.5164 | 0.56569 | 0.6377 | 0.32506 | 3.0969 | 0.32506 | 0.52694 | 0.45898 | 0.42269 | 0.30136 | 0.19408 | 0.18804 | 0.25586 | 0.24457 | 0.37771 | 0.99532 | 0.072111 | 0.12828 | 0.60663 | 0.24833 | 0.75277 | 1.0187 | 0.50761 | 0.37238 | 0.19748 | 0.6047
| Rel.Std.Dev. |  | 3.873% | 3.58% | 5.115% | 2.447% | 17.97% | 2.907% | 4.313% | 3.662% | 2.733% | 2.141% | 1.415% | 1.303% | 2.294% | 1.678% | 2.432% | 6.408% | 0.6462% | 0.9673% | 4.333% | 1.923% | 4.754% | 7.948% | 4.055% | 2.347% | 1.132% | 3.948%
| PDM3 |  | -4.052% | 13.7% | -10.29% | -4.412% | 24.01% | -19.52% | -12.09% | -9.809% | 11.3% | 1.309% | -1.294% | 3.839% | -19.72% | 4.865% | 11.78% | 11.78% | -19.69% | -4.568% | 0.7451% | -7.051% | 13.94% | -7.77% | -9.929% | 14.18% | 25.57% | 10.22%
| Z-Score (Absolute) |  | NA | 0.92 | 0.66 | 0.27 | 1.6 | 1.27 | 0.78 | 0.63 | 0.76 | 0.1 | 0.07 | 0.27 | 1.28 | 0.34 | 0.79 | 0.79 | 1.28 | 0.28 | 0.07 | 0.45 | NA | 0.49 | 0.63 | 0.95 | 1.7 | 0.69
Results from laboratories 1.01 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 627 (Certified Value 1.32 ppm) | [30 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.12; 1.04: 1.13; 1.05: 1.21; 1.06: 1.22; 1.07: 1.23; 1.08: 1.25; 1.09: 1.26; 1.10: 1.30 |  |  |  | ppm
|  |  | 1.01: 4A*MS; 1.02: 4A*MS; 1.03: 4A*MS; 1.04: 4A*MS; 1.05: 4A*OES/MS; 1.06: 4A*MS; 1.07: 4A*MS; 1.08: 4A*OES/MS; 1.09: 4A*OES/MS; 1.10: 4A*MS |  |  |  | 2
2
| 1 | 1 | 1.01: 1.5; 1.02: 1.33; 1.03: 1.35; 1.04: 1.0541; 1.05: 1.3; 1.06: 1.4932; 1.07: 1.2; 1.08: 1.4; 1.09: 1.1; 1.10: 1.34 |  |  |  | 1
| 1 | 2 | 1.01: 1.4; 1.02: 1.29; 1.03: 1.37; 1.04: 1.0363; 1.05: 1.4; 1.06: 1.3697; 1.07: 1.2; 1.08: 1.4; 1.09: 1.1; 1.10: 1.42 |  |  |  | 45
| 1 | 3 | 1.01: 1.5; 1.02: 1.25; 1.03: 1.37; 1.04: 1.0623; 1.05: 1.4; 1.06: 1.4677; 1.07: 1.2; 1.08: 1.5; 1.09: 1.2; 1.10: 1.37 |  |  |  | 16
| 1 | 4 | 1.01: 1.4; 1.02: 1.26; 1.03: 1.44; 1.04: 1.1004; 1.05: 1.3; 1.06: 1.5725; 1.07: 1.2; 1.08: 1.4; 1.09: 1.3; 1.10: 1.32 |  |  |  | 1.3177
| 1 | 5 | 1.01: 1.5; 1.02: 1.27; 1.03: 1.39; 1.04: 1.0333; 1.05: 1.4; 1.06: 1.3824; 1.07: 1.2; 1.08: 1.3; 1.09: 1.1; 1.10: 1.45 |  |  |  | 65
| 1 | 6 | 1.01: 1.5; 1.02: 1.22; 1.03: 1.43; 1.04: 1.0441; 1.05: 1.4; 1.06: 1.5335; 1.07: 1.2; 1.08: 1.5; 1.09: 1.2; 1.10: 1.34
| Mean |  | 1.01: 1.4667; 1.02: 1.27; 1.03: 1.3917; 1.04: 1.0551; 1.05: 1.3667; 1.06: 1.4698; 1.07: 1.2; 1.08: 1.4167; 1.09: 1.1667; 1.10: 1.3733
| Median |  | 1.01: 1.5; 1.02: 1.265; 1.03: 1.38; 1.04: 1.0491; 1.05: 1.4; 1.06: 1.4804; 1.07: 1.2; 1.08: 1.4; 1.09: 1.15; 1.10: 1.355
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.037417; 1.03: 0.036009; 1.04: 0.024709; 1.05: 0.05164; 1.06: 0.081021; 1.07: 0; 1.08: 0.075277; 1.09: 0.08165; 1.10: 0.051251
| Rel.Std.Dev. |  | 1.01: 3.521%; 1.02: 2.946%; 1.03: 2.587%; 1.04: 2.342%; 1.05: 3.779%; 1.06: 5.512%; 1.07: 0%; 1.08: 5.314%; 1.09: 6.999%; 1.10: 3.732%
| PDM3 |  | 1.01: 11.31%; 1.02: -3.617%; 1.03: 5.617%; 1.04: -19.93%; 1.05: 3.719%; 1.06: 11.55%; 1.07: -8.929%; 1.08: 7.514%; 1.09: -11.46%; 1.10: 4.225%
| Z-Score (Absolute) |  | 1.01: 0.66; 1.02: 0.69; 1.03: 0.15; 1.04: 2.16; 1.05: 0.02; 1.06: 0.69; 1.07: 1.17; 1.08: 0.32; 1.09: 1.4; 1.10: 0.02
| Analytical results for Zn in OREAS 627 (Certified Value 10.04 wt.%) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | wt.%
|  |  | 4A*OES | 4A*OES | 4A*AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/AAS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*AAS | 4A*OES | 4A*OES | 4A*OES | 4A*AAS | 4A*AAS | 4A*OES | 4A*OES | 4A*MS | 2
3
| 1 | 1 | 9.6678 | 10.1 | 9.615 | 9.84 | 9.94 | 9.88 | 10.05 | 10.3 | 10.4 | 10.3021 | 10.54 | > 2 | 9.8556 | 10.3037 | 10.3669 | 10.96 | 9.61 | > 1 | 10.0551 | 8.9856 | 9.593 | 9.5604 | 9.97 | 10.25 | 10.33 | 9.901 | 10.01 | > 1 | 10.2435 | 1
| 1 | 2 | 9.503 | 9.93 | 9.769 | 9.9 | 9.95 | 9.76 | 10.1 | 10.25 | 11.2 | 10.1603 | 10.48 | > 2 | 9.8379 | 10.6714 | 10.0563 | 10.91 | 9.52 | > 1 | 10.7154 | 9.067 | 9.445 | 9.5138 | 10.17 | 10.23 | 10.27 | 9.863 | 10.03 | > 1 | 10.607 | 24
| 1 | 3 | 9.452 | 10.1 | 9.874 | 10 | 10.1 | 9.95 | 10.05 | 10.25 | 12.2 | 10.2025 | 10.36 | > 2 | 9.8994 | 10.3778 | 10.0641 | 10.91 | 9.72 | > 1 | 10.7271 | 9.0809 | 9.59 | 9.5068 | 10.14 | 10.21 | 10.38 | 9.887 | 9.93 | > 1 | 9.9688 | 16
| 1 | 4 | 9.5115 | 10.1 | 9.74 | 9.98 | 9.95 | 9.86 | 9.87 | 10.15 | 11.3 | 10.3527 | 10.42 | 10.1367 | 9.9716 | 10.5717 | 10.3806 | 10.91 | 9.75 | > 1 | 10.7206 | 9.0967 | 9.408 | 9.5183 | 10.06 | 10.21 | 10.48 | 9.884 | 10 | > 1 | 9.8883 | 10.0361
| 1 | 5 | 9.656 | 9.98 | 9.768 | 10.1 | 10.1 | 9.9 | 9.97 | 10.35 | 12 | 10.0581 | 10.44 | > 2 | 9.814 | 10.1885 | 10.0643 | 10.99 | 9.7 | > 1 | 10.7218 | 8.8934 | 9.45 | 9.5942 | 10.04 | 10.28 | 10.55 | 9.872 | 9.99 | > 1 | 10.3129 | 66
| 1 | 6 | 9.8515 | 9.97 | 9.741 | 10.2 | 9.98 | 9.85 | 9.91 | 10.25 | 11.5 | 10.122 | 10.49 | > 2 | 9.8934 | 10.4999 | 10.2759 | 10.98 | 9.75 | > 1 | 10.7226 | 9.0265 | 9.62 | 9.5364 | 9.99 | 10.2 | 10.41 | 9.884 | 10.01 | > 1 | 9.7833
| Mean |  | 9.607 | 10.03 | 9.7512 | 10.0033 | 10.0033 | 9.8667 | 9.9917 | 10.2583 | 11.4333 | 10.1996 | 10.455 | 10.1367 | 9.8787 | 10.4355 | 10.2014 | 10.9433 | 9.675 |  | 10.6104 | 9.025 | 9.5177 | 9.5383 | 10.0617 | 10.23 | 10.4033 | 9.8818 | 9.995 |  | 10.134
| Median |  | 9.5838 | 10.04 | 9.7545 | 9.99 | 9.965 | 9.87 | 10.01 | 10.25 | 11.4 | 10.1814 | 10.46 | 10.1367 | 9.8745 | 10.4389 | 10.1701 | 10.935 | 9.71 |  | 10.7212 | 9.0467 | 9.52 | 9.5274 | 10.05 | 10.22 | 10.395 | 9.884 | 10.005 |  | 10.1061
| Std Dev. |  | 0.14822 | 0.078486 | 0.083017 | 0.1311 | 0.07607 | 0.06314 | 0.089981 | 0.066458 | 0.64083 | 0.11094 | 0.06253 |  | 0.05597 | 0.17892 | 0.15733 | 0.037771 | 0.091815 |  | 0.27208 | 0.07603 | 0.093022 | 0.033494 | 0.079854 | 0.030332 | 0.10113 | 0.013075 | 0.034496 |  | 0.30931
| Rel.Std.Dev. |  | 1.543% | 0.7825% | 0.8514% | 1.311% | 0.7604% | 0.6399% | 0.9006% | 0.6478% | 5.605% | 1.088% | 0.5981% |  | 0.5666% | 1.714% | 1.542% | 0.3452% | 0.949% |  | 2.564% | 0.8424% | 0.9774% | 0.3512% | 0.7936% | 0.2965% | 0.9721% | 0.1323% | 0.3451% |  | 3.052%
| PDM3 |  | -4.276% | -0.06042% | -2.839% | -0.3261% | -0.3261% | -1.688% | -0.4424% | 2.215% | 13.92% | 1.63% | 4.174% | 1.003% | -1.568% | 3.98% | 1.647% | 9.04% | -3.598% |  | 5.723% | -10.07% | -5.165% | -4.96% | 0.2551% | 1.932% | 3.659% | -1.537% | -0.4092% |  | 0.9755%
| Z-Score (Absolute) |  | 1.43 | 0 | 0.94 | 0.09 | 0.09 | 0.55 | 0.13 | 0.77 | 4.73 | 0.57 | 1.43 | 0.36 | 0.51 | 1.37 | 0.58 | 3.08 | 1.2 | NA | 1.96 | 3.39 | 1.73 | 1.66 | 0.11 | 0.67 | 1.26 | 0.5 | 0.12 | NA | 0.35
| Analytical results for Zr in OREAS 627 (Certified Value 154 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS |  | 0
0
| 1 | 1 | 147 | 144 | 155.8 | 177.5 | 162.5 | 161 | 140.5 | 159 | 161 | 153.2 | 147.3 | 139.9 | 138.2593 | 145.7743 | 155.4 | 159 | 163.8 | 135 | 146.6 | 157 | 150.12 | 157 | 166 | 159 | 155.3 | 162.3 | 148 | 151 |  | 1
| 1 | 2 | 145 | 147 | 153.8 | 178.5 | 167.5 | 161.5 | 150 | 159 | 167 | 153.6 | 150.7 | 149.6 | 134.6512 | 145.1196 | 154.6 | 162 | 158 | 134 | 146.9 | 165 | 152.43 | 152 | 166 | 160 | 150.3 | 167.5 | 142 | 157 |  | 15
| 1 | 3 | 145 | 148 | 152 | 177.5 | 163 | 152.5 | 146.5 | 157.5 | 163 | 158.7 | 142.8 | 146.5 | 136.1253 | 145.4851 | 160.3 | 162 | 160.2 | 133 | 145.5 | 156 | 157.84 | 151 | 165 | 159 | 156 | 163.9 | 141 | 161 |  | 16
| 1 | 4 | 145 | 155 | 152.5 | 177.5 | 167.5 | 155.5 | 145 | 159 | 161 | 158.7 | 142.2 | 152.1 | 136.7079 | 148.4112 | 151.5 | 161 | 160.3 | 136 | 156 | 148 | 150.68 | 153 | 161 | 159 | 153.4 | 165 | 143 | 154 |  | 154.0845
| 1 | 5 | 144 | 147 | 153.1 | 173 | 149.5 | 156.5 | 145.5 | 154 | 172 | 156.2 | 143.3 | 149.2 | 136.7924 | 147.192 | 151.2 | 163 | 159.2 | 136 | 151.6 | 158 | 151.03 | 153 | 165 | 160 | 155.6 | 163.1 | 143 | 163 |  | 67
| 1 | 6 | 145 | 149 | 155 | 172.5 | 167.5 | 153.5 | 147 | 165 | 162 | 158.6 | 145.4 | 148.9 | 145.912 | 147.8977 | 151.9 | 164 | 161.5 | 135 | 151 | 155 | 155.91 | 159 | 164 | 159 | 153.9 | 164.6 | 145 | 161
| Mean |  | 145.1667 | 148.3333 | 153.7 | 176.0833 | 162.9167 | 156.75 | 145.75 | 158.9167 | 164.3333 | 156.5 | 145.2833 | 147.7 | 138.0747 | 146.6467 | 154.15 | 161.8333 | 160.5 | 134.8333 | 149.6 | 156.5 | 153.0017 | 154.1667 | 164.5 | 159.3333 | 154.0833 | 164.4 | 143.6667 | 157.8333
| Median |  | 145 | 147.5 | 153.45 | 177.5 | 165.25 | 156 | 146 | 159 | 162.5 | 157.4 | 144.35 | 149.05 | 136.7502 | 146.4832 | 153.25 | 162 | 160.25 | 135 | 148.95 | 156.5 | 151.73 | 153 | 165 | 159 | 154.6 | 164.25 | 143 | 159
| Std Dev. |  | 0.98319 | 3.6697 | 1.4697 | 2.6157 | 6.9744 | 3.765 | 3.1105 | 3.5555 | 4.3665 | 2.5877 | 3.2566 | 4.2176 | 4.0122 | 1.3725 | 3.4749 | 1.7224 | 1.998 | 1.169 | 4.0005 | 5.4681 | 3.1552 | 3.1252 | 1.8708 | 0.5164 | 2.1104 | 1.8089 | 2.5033 | 4.6655
| Rel.Std.Dev. |  | 0.6773% | 2.474% | 0.9562% | 1.485% | 4.281% | 2.402% | 2.134% | 2.237% | 2.657% | 1.653% | 2.242% | 2.856% | 2.906% | 0.9359% | 2.254% | 1.064% | 1.245% | 0.867% | 2.674% | 3.494% | 2.062% | 2.027% | 1.137% | 0.3241% | 1.37% | 1.1% | 1.742% | 2.956%
| PDM3 |  | -5.788% | -3.732% | -0.2495% | 14.28% | 5.732% | 1.73% | -5.409% | 3.136% | 6.651% | 1.568% | -5.712% | -4.143% | -10.39% | -4.827% | 0.04252% | 5.029% | 4.164% | -12.49% | -2.91% | 1.568% | -0.7027% | 0.05334% | 6.76% | 3.406% | -0.0007464% | 6.695% | -6.761% | 2.433%
| Z-Score (Absolute) |  | 0.95 | 0.61 | 0.05 | 2.31 | 0.92 | 0.27 | 0.89 | 0.5 | 1.07 | 0.25 | 0.94 | 0.68 | 1.69 | 0.79 | 0 | 0.81 | 0.67 | 2.04 | 0.48 | 0.25 | 0.12 | 0 | 1.09 | 0.55 | 0.01 | 1.08 | 1.11 | 0.39
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C290: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C308: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C326: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C344: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C381: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C399: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C417: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C508: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C636: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C654: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C673: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C691: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C709: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C727: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C745: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C763: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C781: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C799: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C873: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C891: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C910: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C929: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C948: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C966: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1002: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1020: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1038: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1056: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1075: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 627 (Certified Value 62 ppm) | [30 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | 2
| 1 | 1 | 61 | 52.5 | 64.5 | 62.6 | 66.2 | 62.4 | 63.5 | 66.3 | 62.3 | 64.3 | 60.984 | 59.1 | 61.2292 | 50.12 | 58.8346 | 60.3 | 59.8 | 57.2 | 65.01 | 63.9 | 65.8 | 70.092 | 64.57 | 61.9 | 62.3 | > 10 | 62.9 | 58 | 60.63 | 1
| 1 | 2 | 62 | 52.4 | 63.5 | 63.7 | 61.9 | 63.4 | 62.8 | 67.3 | 61.7 | 63.42 | 61.294 | 58.2 | 60.6619 | 60.5 | 59.1982 | 60.21 | 60.1 | 57.7 | 64.53 | 64.1 | 66.1 | 69.8544 | 62.61 | 62.9 | 63.7 | > 10 | 61.69 | 63 | 54.15 | 1
| 1 | 3 | 62 | 52.3 | 64.6 | 62.9 | 68.2 | 64.3 | 61.3 | 66.6 | 61.9 | 61.6 | 60.186 | 59.1 | 61.1138 | 50.2 | 58.8079 | 61.51 | 61.1 | 57 | 63.14 | 63.4 | 64.2 | 70.5456 | 63.88 | 60.9 | 62.7 | > 10 | 62.67 | 63 | 55.44 | 16
| 1 | 4 | 62 | 51 | 64.5 | 64 | 63.3 | 64.2 | 62.1 | 67.7 | 63 | 60.73 | 61.219 | 57.6 | 62.7106 | 59.6174 | 59.7354 | 60.56 | 61.2 | 57 | 64.74 | 63 | 62.8 | 68.634 | 62.77 | 62.3 | 62.5 | > 10 | 62.71 | 63 | 58.77 | 62.0419
| 1 | 5 | 62 | 49.3 | 64.8 | 63.6 | 65.3 | 64.3 | 61.9 | 66 | 64.8 | 61.46 | 61.51 | 58.7 | 61.8021 | 55.97 | 59.9473 | 61.58 | 60.3 | 56.5 | 64.77 | 65 | 64.7 | 67.7916 | 63.02 | 61.8 | 62.7 | > 10 | 60.64 | 62 | 58.68 | 69
| 1 | 6 | 62 | 51.7 | 64.6 | 64.3 | 63.5 | 65.4 | 62.1 | 66 | 61.9 | 58.62 | 61.181 | 58.6 | 61.2941 | 61.32 | 59.0789 | 60.56 | 60.8 | 57.5 | 64.86 | 63.3 | 63.6 | 70.0164 | 63.41 | 62.1 | 63.5 | > 10 | 61.1 | 62 | 60.23
| Mean |  | 61.8333 | 51.5333 | 64.4167 | 63.5167 | 64.7333 | 64 | 62.2833 | 66.65 | 62.6 | 61.6883 | 61.0623 | 58.55 | 61.4686 | 56.2879 | 59.2671 | 60.7867 | 60.55 | 57.15 | 64.5083 | 63.7833 | 64.5333 | 69.489 | 63.3767 | 61.9833 | 62.9 |  | 61.9517 | 61.8333 | 57.9833
| Median |  | 62 | 52 | 64.55 | 63.65 | 64.4 | 64.25 | 62.1 | 66.45 | 62.1 | 61.53 | 61.2 | 58.65 | 61.2616 | 57.7937 | 59.1386 | 60.56 | 60.55 | 57.1 | 64.755 | 63.65 | 64.45 | 69.9354 | 63.215 | 62 | 62.7 |  | 62.18 | 62.5 | 58.725
| Std Dev. |  | 0.40825 | 1.2307 | 0.46224 | 0.64936 | 2.2844 | 1.01 | 0.76529 | 0.7064 | 1.173 | 2.0106 | 0.46191 | 0.57533 | 0.70969 | 5.0864 | 0.47331 | 0.60411 | 0.56833 | 0.42308 | 0.68852 | 0.71949 | 1.2707 | 1.0492 | 0.74317 | 0.65853 | 0.56569 |  | 0.94918 | 1.9408 | 2.6199
| Rel.Std.Dev. |  | 0.6602% | 2.388% | 0.7176% | 1.022% | 3.529% | 1.578% | 1.229% | 1.06% | 1.874% | 3.259% | 0.7565% | 0.9826% | 1.155% | 9.036% | 0.7986% | 0.9938% | 0.9386% | 0.7403% | 1.067% | 1.128% | 1.969% | 1.51% | 1.173% | 1.062% | 0.8993% |  | 1.532% | 3.139% | 4.518%
| PDM3 |  | -0.3361% | -16.94% | 3.828% | 2.377% | 4.338% | 3.156% | 0.3892% | 7.427% | 0.8996% | -0.5698% | -1.579% | -5.628% | -0.924% | -9.274% | -4.472% | -2.023% | -2.405% | -7.885% | 3.976% | 2.807% | 4.016% | 12% | 2.151% | -0.09432% | 1.383% |  | -0.1454% | -0.3361% | -6.542%
| Z-Score (Absolute) |  | 0.05 | 4.18 | 0.98 | 0.62 | 1.11 | 0.81 | 0.13 | 1.88 | 0.25 | 0.11 | 0.36 | 1.37 | 0.2 | 2.28 | 1.08 | 0.47 | 0.57 | 1.93 | 1.02 | 0.73 | 1.03 | 3.01 | 0.56 | 0.01 | 0.37 | NA | 0.01 | 0.05 | 1.6
| Analytical results for Al in OREAS 627 (Certified Value 0.755 wt.%) | [30 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 0.78 | 0.89 | 0.69 | 0.67 | 0.89 | 0.71 | 0.73 | 0.71 | 0.57 | 0.8621 | 0.81 | 0.9815 | 0.74835 | 0.614 | 0.9328 | 0.73 | 0.929 | 0.78 | 0.7338 | 1.27 | 1.2164 | 0.81 | 0.73 | 0.7 | 0.74 | 0.75 | 0.799 | 0.83 |  | 1
| 1 | 2 | 0.75 | 0.7 | 0.75 | 0.69 | 0.8 | 0.71 | 0.71 | 0.75 | 0.57 | 0.8246 | 0.82 | 0.9134 | 0.77663 | 0.674 | 0.93803 | 0.69 | 0.925 | 0.77 | 0.7496 | 1.21 | 1.0326 | 0.81 | 0.71 | 0.71 | 0.72 | 0.73 | 0.8 | 0.75 |  | 0
| 1 | 3 | 0.79 | 0.78 | 0.75 | 0.68 | 0.91 | 0.75 | 0.69 | 0.74 | 0.6 | 0.8041 | 0.8 | 0.939 | 0.77277 | 0.61 | 0.96453 | 0.73 | 0.878 | 0.8 | 0.7523 | 1.21 | 1.1807 | 0.82 | 0.7 | 0.71 | 0.74 | 0.74 | 0.801 | 0.77 |  | 16
| 1 | 4 | 0.82 | 0.85 | 0.77 | 0.68 | 0.83 | 0.76 | 0.71 | 0.75 | 0.57 | 0.7659 | 0.79 | 0.934 | 0.78766 | 0.654 | 0.91601 | 0.67 | 0.934 | 0.81 | 0.7486 | 1.26 | 1.114 | 0.81 | 0.72 | 0.7 | 0.74 | 0.76 | 0.807 | 0.8 |  | 0.75513
| 1 | 5 | 0.82 | 0.68 | 0.78 | 0.69 | 0.83 | 0.75 | 0.71 | 0.74 | 0.58 | 0.7723 | 0.85 | 0.9463 | 0.78384 | 0.628 | 0.90019 | 0.78 | 0.922 | 0.82 | 0.7506 | 1.29 | 1.1107 | 0.81 | 0.71 | 0.7 | 0.75 | 0.73 | 0.804 | 0.73 |  | 70
| 1 | 6 | 0.83 | 0.77 | 0.73 | 0.72 | 0.82 | 0.78 | 0.7 | 0.71 | 0.58 | 0.6826 | 0.82 | 0.9966 | 0.76027 | 0.663 | 0.94382 | 0.77 | 0.937 | 0.78 | 0.7518 | 1.27 | 1.1285 | 0.79 | 0.72 | 0.71 | 0.75 | 0.75 | 0.785 | 0.79
| Mean |  | 0.79833 | 0.77833 | 0.745 | 0.68833 | 0.84667 | 0.74333 | 0.70833 | 0.73333 | 0.57833 | 0.78527 | 0.815 | 0.9518 | 0.77159 | 0.6405 | 0.93256 | 0.72833 | 0.92083 | 0.79333 | 0.74778 | 1.2517 | 1.1305 | 0.80833 | 0.715 | 0.705 | 0.74 | 0.74333 | 0.79933 | 0.77833
| Median |  | 0.805 | 0.775 | 0.75 | 0.685 | 0.83 | 0.75 | 0.71 | 0.74 | 0.575 | 0.7882 | 0.815 | 0.94265 | 0.7747 | 0.641 | 0.93542 | 0.73 | 0.927 | 0.79 | 0.7501 | 1.265 | 1.1213 | 0.81 | 0.715 | 0.705 | 0.74 | 0.745 | 0.8005 | 0.78
| Std Dev. |  | 0.030605 | 0.081833 | 0.032094 | 0.017224 | 0.043205 | 0.028048 | 0.013292 | 0.018619 | 0.01169 | 0.061464 | 0.020736 | 0.031224 | 0.014857 | 0.026831 | 0.022362 | 0.043089 | 0.021702 | 0.019664 | 0.0069852 | 0.033714 | 0.063485 | 0.0098319 | 0.010488 | 0.0054772 | 0.010954 | 0.012111 | 0.007607 | 0.036009
| Rel.Std.Dev. |  | 3.834% | 10.51% | 4.308% | 2.502% | 5.103% | 3.773% | 1.876% | 2.539% | 2.021% | 7.827% | 2.544% | 3.28% | 1.925% | 4.189% | 2.398% | 5.916% | 2.357% | 2.479% | 0.9341% | 2.694% | 5.616% | 1.216% | 1.467% | 0.7769% | 1.48% | 1.629% | 0.9517% | 4.626%
| PDM3 |  | 5.722% | 3.073% | -1.341% | -8.845% | 12.12% | -1.562% | -6.197% | -2.886% | -23.41% | 3.991% | 7.929% | 26.04% | 2.18% | -15.18% | 23.5% | -3.548% | 21.94% | 5.06% | -0.9726% | 65.76% | 49.71% | 7.046% | -5.314% | -6.638% | -2.003% | -1.562% | 5.854% | 3.073%
| Z-Score (Absolute) |  | 0.39 | 0.06 | 0.49 | 1.43 | 1.18 | 0.52 | 1.1 | 0.69 | 3.25 | 0.17 | 0.66 | 2.92 | 0.06 | 2.22 | 2.6 | 0.77 | 2.41 | 0.3 | 0.45 | 7.87 | 5.87 | 0.55 | 0.99 | 1.16 | 0.58 | 0.52 | 0.4 | 0.06
| Analytical results for As in OREAS 627 (Certified Value 299 ppm) | [30 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | 0
| 1 | 1 | 291 | 291 | 313 | 300 | 329 | 301 | 305 | 306 | 283 | 314 | 324.5 | 301 | 281.2489 | 208.337 | 290.8791 | 294.8 | 299 | 253.9 | 275 | 301.1 | 326 | 274.84 | 316 | 293 | 291 | 282 | 296 | 288 | 343.7 | 1
| 1 | 2 | 279 | 286 | 311.8 | 307 | 304 | 308 | 300 | 312 | 285 | 314 | 318.6 | 296 | 295.7838 | 254.81 | 288.6803 | 293.9 | 296 | 256.2 | 265 | 309 | 317 | 277.42 | 309 | 295 | 295 | 290 | 301.8 | 301 | 308.1 | 21
| 1 | 3 | 292 | 285 | 320.3 | 301 | 339 | 312 | 300 | 309 | 291 | 313 | 314.4 | 300 | 289.6294 | 206.155 | 285.8994 | 293.2 | 300 | 258.7 | 258 | 298.5 | 315 | 281.01 | 305 | 288 | 295 | 289 | 304.1 | 302 | 317.2 | 16
| 1 | 4 | 291 | 281 | 315 | 305 | 311 | 316 | 299 | 313 | 287 | 318 | 323.3 | 293 | 293.7352 | 250.538 | 294.0489 | 293.7 | 301 | 256.6 | 264 | 302.6 | 310 | 275.36 | 297 | 293 | 293 | 290 | 306.7 | 297 | 335.1 | 298.8152
| 1 | 5 | 291 | 275 | 317 | 304 | 323 | 307 | 302 | 306 | 296 | 310 | 320.9 | 299 | 290.2496 | 227.671 | 295.3871 | 293.8 | 294 | 259.6 | 266 | 298.5 | 321 | 279.56 | 309 | 289 | 291 | 292 | 288 | 300 | 350.5 | 71
| 1 | 6 | 291 | 286 | 318.1 | 306 | 315 | 309 | 299 | 305 | 288 | 299 | 322.1 | 297 | 292.4704 | 240.29 | 289.2318 | 297.4 | 299 | 257.4 | 272 | 299 | 313 | 281.04 | 303 | 292 | 294 | 287 | 300.2 | 297 | 341.9
| Mean |  | 289.1667 | 284 | 315.8667 | 303.8333 | 320.1667 | 308.8333 | 300.8333 | 308.5 | 288.3333 | 311.3333 | 320.6333 | 297.6667 | 290.5195 | 231.3002 | 290.6878 | 294.4667 | 298.1667 | 257.0667 | 266.6667 | 301.45 | 317 | 278.205 | 306.5 | 291.6667 | 293.1667 | 288.3333 | 299.4667 | 297.5 | 332.75
| Median |  | 291 | 285.5 | 316 | 304.5 | 319 | 308.5 | 300 | 307.5 | 287.5 | 313.5 | 321.5 | 298 | 291.36 | 233.9805 | 290.0555 | 293.85 | 299 | 257 | 265.5 | 300.05 | 316 | 278.49 | 307 | 292.5 | 293.5 | 289.5 | 301 | 298.5 | 338.5
| Std Dev. |  | 4.9967 | 5.4406 | 3.2073 | 2.7869 | 12.7502 | 5.0365 | 2.3166 | 3.3912 | 4.6332 | 6.5625 | 3.6681 | 2.9439 | 5.0729 | 20.8633 | 3.5354 | 1.528 | 2.6394 | 2.0096 | 6.0553 | 4.0461 | 5.7619 | 2.7486 | 6.442 | 2.6583 | 1.8348 | 3.5024 | 6.6824 | 5.0892 | 16.5749
| Rel.Std.Dev. |  | 1.728% | 1.916% | 1.015% | 0.9172% | 3.982% | 1.631% | 0.7701% | 1.099% | 1.607% | 2.108% | 1.144% | 0.989% | 1.746% | 9.02% | 1.216% | 0.5189% | 0.8852% | 0.7818% | 2.271% | 1.342% | 1.818% | 0.988% | 2.102% | 0.9114% | 0.6259% | 1.215% | 2.231% | 1.711% | 4.981%
| PDM3 |  | -3.229% | -4.958% | 5.706% | 1.679% | 7.145% | 3.353% | 0.6754% | 3.241% | -3.508% | 4.189% | 7.302% | -0.3844% | -2.776% | -22.59% | -2.72% | -1.455% | -0.217% | -13.97% | -10.76% | 0.8817% | 6.086% | -6.897% | 2.572% | -2.392% | -1.89% | -3.508% | 0.218% | -0.4401% | 11.36%
| Z-Score (Absolute) |  | 0.61 | 0.99 | 1.31 | 0.45 | 1.63 | 0.81 | 0.23 | 0.78 | 0.67 | 0.99 | 1.66 | 0 | 0.52 | 4.79 | 0.5 | 0.23 | 0.04 | 2.93 | 2.24 | 0.27 | 1.4 | 1.41 | 0.64 | 0.43 | 0.33 | 0.67 | 0.13 | 0.01 | 2.53
| Analytical results for B in OREAS 627 (Certified Value < 10 ppm) | [30 empty] | Certified
B | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.13 | 1.14 | 1.15 | 1.26 | 1.27 | [14 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [14 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 15g | 0.2g | 0.5g | 01g | 0.25g | 0.5g | [14 empty] | 1
| 1 | 1 | 28 | 4 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 63.648 | < 10 | < 10 | 10 | [14 empty] | 1
| 1 | 2 | 26 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 60.813 | < 10 | < 10 | 11 | [14 empty] | 3
| 1 | 3 | 29 | 4 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 62.898 | < 10 | < 10 | 10 | [14 empty] | 16
| 1 | 4 | 30 | 4 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 59.348 | < 10 | < 10 | 10 | [14 empty] | < 10
| 1 | 5 | 28 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 61.748 | < 10 | < 10 | 12 | [14 empty] | 72
| 1 | 6 | 29 | 3 | < 20 | < 10 | < 10 | < 10 | < 10 | < 10 | < 5 | 2 | < 2 | 66.768 | < 10 | < 10 | 13
| Mean |  | 1.01: 28.3333; 1.02: 3.5; 1.11: 2; 1.14: 62.5372; 1.27: 11
| Median |  | 1.01: 28.5; 1.02: 3.5; 1.11: 2; 1.14: 62.323; 1.27: 10.5
| Std Dev. |  | 1.01: 1.3663; 1.02: 0.54772; 1.11: 0; 1.14: 2.5702; 1.27: 1.2649
| Rel.Std.Dev. |  | 1.01: 4.822%; 1.02: 15.65%; 1.11: 0%; 1.14: 4.11%; 1.27: 11.5%
PDM3
| Z-Score (Absolute) |  | 10.49 | 0.67 | 2.25 | 0 | 0 | 0 | 0 | 0 | 1.12 | 1.35 | 1.8 | 25.87 | 0 | 0 | 2.7
| Analytical results for Ba in OREAS 627 (Indicative Value 42.5 ppm) | [30 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 1
| 1 | 1 | < 10 | 2.6 | 11 | 10 | 60 | 10 | 20 | 10 | 99 | 13 | 11.4 | 135 | 34.6529 | 75.84 | 8.7237 | 2028 | 98 | 4 | 406 | 44 | 23 | 38 | 31 | 23 | 120 | 215 |  |  |  | 1
| 1 | 2 | < 10 | 2.1 | 12 | 10 | 40 | 10 | 10 | 10 | 98 | 10 | 11 | 87 | 35.975 | 79.75 | 11.8797 | 1994 | 104 | 4 | 384 | 39 | 24 | 37 | 30 | 24 | 126 | 226 |  |  |  | 4
| 1 | 3 | < 10 | 1.7 | 11 | 10 | 50 | 10 | 10 | 10 | 104 | 7 | 10.4 | 130 | 33.8552 | 73.08 | 13.8717 | 2006 | 92 | 4 | 377 | 40 | 21 | 37 | 32 | 23 | 106 | 228 |  |  |  | 16
| 1 | 4 | < 10 | 2.5 | 12 | 10 | 50 | 10 | 20 | 10 | 100 | 9 | 10.1 | 103 | 33.4631 | 82.77 | 11.2626 | 1998 | 97 | 4 | 392 | 42 | 24 | 37 | 31 | 23 | 116 | 234 |  |  |  | 42.5069
| 1 | 5 | < 10 | 3.8 | 11 | 10 | 50 | 10 | 20 | 10 | 103 | 8 | 10.8 | 125 | 37.8868 | 78.128 | 14.6288 | 2028 | 98 | 4 | 396 | 39 | 22 | 37 | 33 | 25 | 103 | 69 |  |  |  | 12
| 1 | 6 | < 10 | 1.6 | 11 | 10 | 50 | 10 | 10 | 10 | 103 | 8 | 9.4 | 118 | 35.0238 | 73.09 | 8.9109 | 2048 | 100 | 5 | 383 | 34 | 22 | 38 | 32 | 28 | 124 | 205
| Mean |  |  | 2.3833 | 11.3333 | 10 | 50 | 10 | 15 | 10 | 101.1667 | 9.1667 | 10.5167 | 116.3333 | 35.1428 | 77.1097 | 11.5462 | 2017 | 98.1667 | 4.1667 | 389.6667 | 39.6667 | 22.6667 | 37.3333 | 31.5 | 24.3333 | 115.8333 | 196.1667
| Median |  |  | 2.3 | 11 | 10 | 50 | 10 | 15 | 10 | 101.5 | 8.5 | 10.6 | 121.5 | 34.8384 | 76.984 | 11.5711 | 2017 | 98 | 4 | 388 | 39.5 | 22.5 | 37 | 31.5 | 23.5 | 118 | 220.5
| Std Dev. |  |  | 0.80353 | 0.5164 | 0 | 6.3246 | 0 | 5.4772 | 0 | 2.4833 | 2.137 | 0.7111 | 18.1732 | 1.6097 | 3.8485 | 2.45 | 21.0428 | 3.92 | 0.40825 | 10.4817 | 3.3862 | 1.2111 | 0.5164 | 1.0488 | 1.9664 | 9.4745 | 63.1488
| Rel.Std.Dev. |  |  | 33.71% | 4.556% | 0% | 12.65% | 0% | 36.51% | 0% | 2.455% | 23.31% | 6.762% | 15.62% | 4.58% | 4.991% | 21.22% | 1.043% | 3.993% | 9.798% | 2.69% | 8.537% | 5.343% | 1.383% | 3.33% | 8.081% | 8.179% | 32.19%
| PDM3 |  |  | -94.39% | -73.34% | -76.47% | 17.63% | -76.47% | -64.71% | -76.47% | 138% | -78.43% | -75.26% | 173.7% | -17.32% | 81.4% | -72.84% | 4645% | 130.9% | -90.2% | 816.7% | -6.682% | -46.68% | -12.17% | -25.89% | -42.75% | 172.5% | 361.5%
| Z-Score (Absolute) |  | 0.78 | 0.85 | 0.62 | NA | 0.39 | NA | NA | 0.65 | 1.71 | 0.67 | 0.64 | 2.11 | 0 | 1.09 | 0.61 | 51.45 | 1.64 | 0.8 | 9.2 | 0.12 | 0.32 | 0.06 | 0.09 | 0.28 | 2.09 | 4.18
Results from laboratories 1.04, 1.06 and 1.07 were removed due to their 10 ppm reading resolution.
| Analytical results for Be in OREAS 627 (Certified Value 0.47 ppm) | [30 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [7 empty] | 2
| 1 | 1 | < 5 | 0.5 | 0.39 | 0.47 | 0.42 | 0.42 | 0.41 | 0.36 | 0.49 | < 0.5 | < 0.5 | < 2 | 0.54 | < 1 | < 5 | 0.6 | 0.5 | 0.5 | 0.5 | 0.43 | < 0.5 | 0.68 | [7 empty] | 1
| 1 | 2 | < 5 | 0.4 | 0.4 | 0.43 | 0.43 | 0.42 | 0.42 | 0.38 | 0.49 | < 0.5 | 0.5083 | < 2 | 0.57 | < 1 | < 5 | 0.6 | 0.5 | 0.5 | 0.5 | 0.43 | < 0.5 | 0.48 | [7 empty] | 14
| 1 | 3 | < 5 | 0.5 | 0.39 | 0.48 | 0.44 | 0.41 | 0.41 | 0.41 | 0.49 | < 0.5 | 0.51455 | < 2 | 0.54 | < 1 | < 5 | 0.5 | 0.5 | 0.5 | 0.5 | 0.43 | < 0.5 | 0.5 | [7 empty] | 16
| 1 | 4 | < 5 | 0.5 | 0.4 | 0.45 | 0.44 | 0.41 | 0.42 | 0.28 | 0.51 | < 0.5 | < 0.5 | < 2 | 0.56 | < 1 | < 5 | 0.6 | 0.5 | 0.5 | 0.5 | 0.45 | < 0.5 | 0.64 | [7 empty] | 0.47359
| 1 | 5 | < 5 | 0.4 | 0.4 | 0.45 | 0.43 | 0.42 | 0.42 | 0.38 | 0.48 | < 0.5 | < 0.5 | < 2 | 0.59 | < 1 | < 5 | 0.6 | 0.5 | 0.5 | 0.5 | 0.43 | < 0.5 | 0.43 | [7 empty] | 73
| 1 | 6 | < 5 | 0.4 | 0.41 | 0.45 | 0.45 | 0.41 | 0.41 | 0.4 | 0.46 | < 0.5 | < 0.5 | < 2 | 0.57 | < 1 | < 5 | 0.6 | 0.5 | 0.5 | 0.5 | 0.46 | < 0.5 | 0.7
| Mean |  |  | 0.45 | 0.39833 | 0.455 | 0.435 | 0.415 | 0.415 | 0.36833 | 0.48667 |  | 0.51142 |  | 0.56167 |  |  | 0.58333 | 0.5 | 0.5 | 0.5 | 0.43833 |  | 0.57167
| Median |  |  | 0.45 | 0.4 | 0.45 | 0.435 | 0.415 | 0.415 | 0.38 | 0.49 |  | 0.51142 |  | 0.565 |  |  | 0.6 | 0.5 | 0.5 | 0.5 | 0.43 |  | 0.57
| Std Dev. |  |  | 0.054772 | 0.0075277 | 0.017607 | 0.010488 | 0.0054772 | 0.0054772 | 0.046655 | 0.01633 |  | 0.0044174 |  | 0.019408 |  |  | 0.040825 | 0 | 0 | 0 | 0.013292 |  | 0.11531
| Rel.Std.Dev. |  |  | 12.17% | 1.89% | 3.87% | 2.411% | 1.32% | 1.32% | 12.67% | 3.355% |  | 0.8638% |  | 3.455% |  |  | 6.999% | 0% | 0% | 0% | 3.032% |  | 20.17%
| PDM3 |  |  | -4.981% | -15.89% | -3.925% | -8.148% | -12.37% | -12.37% | -22.23% | 2.761% |  | 7.989% |  | 18.6% |  |  | 23.17% | 5.577% | 5.577% | 5.577% | -7.444% |  | 20.71%
| Z-Score (Absolute) |  | 21.32 | 0.22 | 0.76 | 0.17 | 0.38 | 0.59 | 0.59 | 1.08 | 0.17 | 2.32 | 1.4 | 5.56 | 0.95 | 0.31 | 21.32 | 1.18 | 0.31 | 0.31 | 0.31 | 0.34 | 2.32 | 1.06
| Analytical results for Bi in OREAS 627 (Certified Value 44.2 ppm) | [30 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 2
| 1 | 1 | 45 | 41.3 | 42 | 45.2 | 46 | 47.6 | 45.9 | 44.5 | 45.2 | 46.58 | 45.16 | 44 | 38.1916 | 52.24 | 44.5972 | 26.8 | 44.9 | 37.7 | 37.53 | 44.86 | 44.23 | 49.12 | 50 | 39.2 | 42.74 | 42.98 | 47 | 49.87 |  | 1
| 1 | 2 | 41 | 41.1 | 41.1 | 45.5 | 43.1 | 50.2 | 46.3 | 44.3 | 45.6 | 46.11 | 44.08 | 41 | 40.2725 | 54.49 | 44.4331 | 22.7 | 45.4 | 36.2 | 37.26 | 45.06 | 44.3 | 46.23 | 52 | 40.8 | 41.85 | 42.38 | 57 | 46.95 |  | 35
| 1 | 3 | 45 | 41.2 | 45.5 | 44.7 | 46.5 | 47.1 | 45 | 43.3 | 46.3 | 45.34 | 42.69 | 41 | 40.4498 | 52.08 | 44.9945 | 25.3 | 45.4 | 37.4 | 37.04 | 44.37 | 43.56 | 45.32 | 52 | 41 | 43.78 | 42.47 | 40 | 47.34 |  | 16
| 1 | 4 | 44 | 41.3 | 44.1 | 44.6 | 43.6 | 50.6 | 45.5 | 44.6 | 46 | 45.44 | 43.6 | 40 | 40.9142 | 51.99 | 44.4317 | 22.8 | 45.8 | 36.7 | 37.33 | 44.42 | 42.88 | 45.74 | 50 | 40.9 | 43.5 | 43.56 | 45 | 48.76 |  | 44.1561
| 1 | 5 | 44 | 40.5 | 45.3 | 46.7 | 45.2 | 47.6 | 45 | 44 | 46.4 | 45.53 | 45.1 | 41 | 41.1663 | 51.196 | 44.9449 | 13.9 | 45.1 | 36.9 | 37.36 | 45.57 | 44.65 | 46.21 | 52 | 40.1 | 41.86 | 41.02 | 28 | 46.16 |  | 74
| 1 | 6 | 44 | 40.6 | 42 | 44 | 43.1 | 50.3 | 45 | 44.1 | 45.3 | 43.25 | 42.65 | 40 | 41.0452 | 52.88 | 44.3525 | 30.9 | 45.6 | 37.3 | 37.38 | 44.67 | 43.3 | 46.99 | 52 | 41.4 | 42.85 | 41.32 | 44 | 48.29
| Mean |  | 43.8333 | 41 | 43.3333 | 45.1167 | 44.5833 | 48.9 | 45.45 | 44.1333 | 45.8 | 45.375 | 43.88 | 41.1667 | 40.3399 | 52.4793 | 44.6257 | 23.7333 | 45.3667 | 37.0333 | 37.3167 | 44.825 | 43.82 | 46.6017 | 51.3333 | 40.5667 | 42.7633 | 42.2883 | 43.5 | 47.895
| Median |  | 44 | 41.15 | 43.05 | 44.95 | 44.4 | 48.9 | 45.25 | 44.2 | 45.8 | 45.485 | 43.84 | 41 | 40.682 | 52.16 | 44.5151 | 24.05 | 45.4 | 37.1 | 37.345 | 44.765 | 43.895 | 46.22 | 52 | 40.85 | 42.795 | 42.425 | 44.5 | 47.815
| Std Dev. |  | 1.472 | 0.35777 | 1.8811 | 0.93256 | 1.5118 | 1.6223 | 0.55408 | 0.46762 | 0.5099 | 1.1434 | 1.1112 | 1.472 | 1.1083 | 1.1231 | 0.27861 | 5.6895 | 0.3266 | 0.54283 | 0.16207 | 0.44868 | 0.67956 | 1.3542 | 1.0328 | 0.79162 | 0.80431 | 0.96781 | 9.4816 | 1.3428
| Rel.Std.Dev. |  | 3.358% | 0.8726% | 4.341% | 2.067% | 3.391% | 3.318% | 1.219% | 1.06% | 1.113% | 2.52% | 2.532% | 3.576% | 2.747% | 2.14% | 0.6243% | 23.97% | 0.7199% | 1.466% | 0.4343% | 1.001% | 1.551% | 2.906% | 2.012% | 1.951% | 1.881% | 2.289% | 21.8% | 2.804%
| PDM3 |  | -0.7309% | -7.148% | -1.863% | 2.175% | 0.9676% | 10.74% | 2.93% | -0.05152% | 3.723% | 2.76% | -0.6252% | -6.77% | -8.642% | 18.85% | 1.063% | -46.25% | 2.742% | -16.13% | -15.49% | 1.515% | -0.7611% | 5.538% | 16.25% | -8.129% | -3.154% | -4.23% | -1.486% | 8.468%
| Z-Score (Absolute) |  | 0.07 | 1.28 | 0.29 | 0.47 | 0.25 | 2.09 | 0.62 | 0.05 | 0.76 | 0.58 | 0.05 | 1.21 | 1.56 | 3.61 | 0.26 | 8.65 | 0.58 | 2.97 | 2.85 | 0.35 | 0.08 | 1.11 | 3.13 | 1.47 | 0.53 | 0.73 | 0.22 | 1.66
| Analytical results for Ca in OREAS 627 (Certified Value 0.645 wt.%) | [30 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 0.66 | 0.57 | 0.43 | 0.65 | 0.67 | 0.65 | 0.64 | 0.68 | 0.6 | 0.65 | 0.63 | 0.66 | 0.63837 | 0.61 | 0.67772 | 0.65 | 0.64 | 0.65 | 0.7357 | 0.6294 | 0.65 | 0.66 | 0.71 | 0.64 | 0.66 | 0.65 | 0.643 | 0.72 |  | 1
| 1 | 2 | 0.67 | 0.58 | 0.4 | 0.66 | 0.63 | 0.66 | 0.65 | 0.67 | 0.61 | 0.64 | 0.63 | 0.66 | 0.63161 | 0.581 | 0.66173 | 0.65 | 0.64 | 0.61 | 0.7345 | 0.6289 | 0.65 | 0.63 | 0.72 | 0.65 | 0.66 | 0.63 | 0.613 | 0.59 |  | 0
| 1 | 3 | 0.68 | 0.59 | 0.36 | 0.66 | 0.69 | 0.66 | 0.65 | 0.66 | 0.64 | 0.62 | 0.61 | 0.67 | 0.63196 | 0.848 | 0.67127 | 0.66 | 0.67 | 0.63 | 0.7339 | 0.6207 | 0.63 | 0.64 | 0.7 | 0.65 | 0.67 | 0.67 | 0.62 | 0.59 |  | 16
| 1 | 4 | 0.64 | 0.58 | 0.37 | 0.66 | 0.68 | 0.66 | 0.65 | 0.67 | 0.62 | 0.63 | 0.6 | 0.68 | 0.64424 | 0.771 | 0.69382 | 0.63 | 0.66 | 0.64 | 0.7351 | 0.6555 | 0.63 | 0.63 | 0.72 | 0.66 | 0.64 | 0.65 | 0.625 | 0.7 |  | 0.64541
| 1 | 5 | 0.65 | 0.57 | 0.36 | 0.66 | 0.67 | 0.64 | 0.66 | 0.66 | 0.62 | 0.63 | 0.62 | 0.67 | 0.64284 | 0.933 | 0.65994 | 0.63 | 0.64 | 0.64 | 0.7352 | 0.6472 | 0.64 | 0.63 | 0.71 | 0.66 | 0.65 | 0.66 | 0.617 | 0.59 |  | 75
| 1 | 6 | 0.65 | 0.58 | 0.39 | 0.66 | 0.66 | 0.65 | 0.65 | 0.67 | 0.59 | 0.59 | 0.6 | 0.68 | 0.65449 | 0.574 | 0.68421 | 0.64 | 0.65 | 0.65 | 0.7366 | 0.6373 | 0.63 | 0.65 | 0.7 | 0.68 | 0.64 | 0.65 | 0.632 | 0.73
| Mean |  | 0.65833 | 0.57833 | 0.385 | 0.65833 | 0.66667 | 0.65333 | 0.65 | 0.66833 | 0.61333 | 0.62667 | 0.615 | 0.67 | 0.64059 | 0.7195 | 0.67478 | 0.64333 | 0.65 | 0.63667 | 0.73517 | 0.6365 | 0.63833 | 0.64 | 0.71 | 0.65667 | 0.65333 | 0.65167 | 0.625 | 0.65333
| Median |  | 0.655 | 0.58 | 0.38 | 0.66 | 0.67 | 0.655 | 0.65 | 0.67 | 0.615 | 0.63 | 0.615 | 0.67 | 0.64061 | 0.6905 | 0.6745 | 0.645 | 0.645 | 0.64 | 0.73515 | 0.63335 | 0.635 | 0.635 | 0.71 | 0.655 | 0.655 | 0.65 | 0.6225 | 0.645
| Std Dev. |  | 0.01472 | 0.0075277 | 0.027386 | 0.0040825 | 0.020656 | 0.008165 | 0.0063246 | 0.0075277 | 0.017512 | 0.020656 | 0.013784 | 0.0089443 | 0.0086204 | 0.15303 | 0.013135 | 0.012111 | 0.012649 | 0.015055 | 0.00093737 | 0.012915 | 0.0098319 | 0.012649 | 0.0089443 | 0.013663 | 0.012111 | 0.013292 | 0.011009 | 0.070048
| Rel.Std.Dev. |  | 2.236% | 1.302% | 7.113% | 0.6201% | 3.098% | 1.25% | 0.973% | 1.126% | 2.855% | 3.296% | 2.241% | 1.335% | 1.346% | 21.27% | 1.947% | 1.882% | 1.946% | 2.365% | 0.1275% | 2.029% | 1.54% | 1.976% | 1.26% | 2.081% | 1.854% | 2.04% | 1.761% | 10.72%
| PDM3 |  | 2.003% | -10.39% | -40.35% | 2.003% | 3.294% | 1.228% | 0.7116% | 3.552% | -4.97% | -2.904% | -4.711% | 3.81% | -0.7471% | 11.48% | 4.551% | -0.3213% | 0.7116% | -1.354% | 13.91% | -1.38% | -1.096% | -0.8378% | 10.01% | 1.745% | 1.228% | 0.9698% | -3.162% | 1.228%
| Z-Score (Absolute) |  | 0.35 | 3.43 | 12.58 | 0.35 | 0.75 | 0.12 | 0.04 | 0.83 | 1.77 | 1.14 | 1.7 | 0.91 | 0.48 | 3.25 | 1.13 | 0.35 | 0.04 | 0.67 | 3.99 | 0.68 | 0.59 | 0.51 | 2.8 | 0.28 | 0.12 | 0.04 | 1.22 | 0.12
| Analytical results for Cd in OREAS 627 (Certified Value 251 ppm) | [30 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 0
| 1 | 1 | 263 | 227 | 227.6 | 242 | 261 | 247 | 243 | 245 | 252 | 266.07 | 249.71 | 254.2 | 252.2226 | 278.397 | 240.9204 | 264.6 | 250 | 195.5 | 230 | 264 | 264 | 250.6 | 254 | 254.33 | 232.95 | 260 | 217 |  |  | 1
| 1 | 2 | 257 | 227 | 226.9 | 248 | 244 | 251 | 239 | 247 | 251 | 261.3 | 252.59 | 245.5 | 255.9567 | 292.873 | 237.5288 | 263.9 | 252 | 193.7 | 234 | 260 | 261 | 252.8 | 268 | 256.24 | 227.77 | 263 | 215 |  |  | 0
| 1 | 3 | 265 | 226 | 233.6 | 243 | 272 | 253 | 239 | 245 | 260 | 253.05 | 244.47 | 250.6 | 256.2743 | 274.007 | 235.9926 | 264 | 247 | 196.6 | 229 | 258 | 258 | 248.1 | 266 | 256.45 | 231.29 | 265 | 212 |  |  | 16
| 1 | 4 | 267 | 224 | 231 | 244 | 248 | 256 | 240 | 248 | 255 | 251.54 | 247.3 | 243.6 | 257.6764 | 281.488 | 243.1469 | 264.6 | 255 | 194.7 | 235 | 254 | 250 | 250.9 | 265 | 260.1 | 234.02 | 265 | 215 |  |  | 250.742
| 1 | 5 | 268 | 220 | 237 | 245 | 258 | 245 | 243 | 244 | 261 | 258.69 | 252.1 | 249.1 | 256.0272 | 272.795 | 243.1403 | 262.3 | 250 | 198.3 | 231 | 265 | 263 | 248.7 | 260 | 258.75 | 226.62 | 265 | 211 |  |  | 76
| 1 | 6 | 268 | 224 | 236.4 | 246 | 251 | 252 | 239 | 244 | 256 | 244.21 | 243.01 | 247.4 | 257.8209 | 283.885 | 239.4725 | 265.2 | 250 | 196.9 | 232 | 259 | 250 | 251.4 | 261 | 260.18 | 226.69 | 261 | 215
| Mean |  | 264.6667 | 224.6667 | 232.0833 | 244.6667 | 255.6667 | 250.6667 | 240.5 | 245.5 | 255.8333 | 255.81 | 248.1967 | 248.4 | 255.9964 | 280.5742 | 240.0336 | 264.1 | 250.6667 | 195.95 | 231.8333 | 260 | 257.6667 | 250.4167 | 262.3333 | 257.675 | 229.89 | 263.1667 | 214.1667
| Median |  | 266 | 225 | 232.3 | 244.5 | 254.5 | 251.5 | 239.5 | 245 | 255.5 | 255.87 | 248.505 | 248.25 | 256.1508 | 279.9425 | 240.1965 | 264.3 | 250 | 196.05 | 231.5 | 259.5 | 259.5 | 250.75 | 263 | 257.6 | 229.53 | 264 | 215
| Std Dev. |  | 4.2269 | 2.6583 | 4.3185 | 2.1602 | 10.1719 | 4.0332 | 1.9748 | 1.6432 | 4.0702 | 7.7989 | 3.9603 | 3.7794 | 2.0235 | 7.3681 | 2.9343 | 1 | 2.6583 | 1.6538 | 2.3166 | 4.0497 | 6.2823 | 1.7452 | 5.0859 | 2.369 | 3.2805 | 2.2286 | 2.2286
| Rel.Std.Dev. |  | 1.597% | 1.183% | 1.861% | 0.8829% | 3.979% | 1.609% | 0.8211% | 0.6693% | 1.591% | 3.049% | 1.596% | 1.522% | 0.7904% | 2.626% | 1.222% | 0.3786% | 1.061% | 0.844% | 0.9993% | 1.558% | 2.438% | 0.6969% | 1.939% | 0.9194% | 1.427% | 0.8468% | 1.041%
| PDM3 |  | 5.553% | -10.4% | -7.441% | -2.423% | 1.964% | -0.03003% | -4.085% | -2.091% | 2.031% | 2.021% | -1.015% | -0.934% | 2.096% | 11.9% | -4.271% | 5.327% | -0.03003% | -21.85% | -7.541% | 3.692% | 2.762% | -0.1297% | 4.623% | 2.765% | -8.316% | 4.955% | -14.59%
| Z-Score (Absolute) |  | 1.01 | 1.88 | 1.34 | 0.43 | 0.36 | 0 | 0.73 | 0.37 | 0.37 | 0.37 | 0.18 | 0.16 | 0.39 | 2.16 | 0.77 | 0.97 | 0 | 3.95 | 1.36 | 0.67 | 0.51 | 0.02 | 0.84 | 0.51 | 1.5 | 0.9 | 2.64
| Analytical results for Ce in OREAS 627 (Certified Value 35.2 ppm) | [30 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [9 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [9 empty] | 1
| 1 | 1 | 28 | 30.8 | 35 | 31.7 | 33.3 | 32 | 30.7 | 36.223 | 51 | 35.9425 | 42.3788 | 50.5 | 27 | 40.57 | 53.8 | 36.54 | 38.8 | 30.18 | 36.18 | 42.64 | [9 empty] | 1
| 1 | 2 | 26.6 | 30.6 | 32.1 | 33.3 | 32.5 | 33.3 | 31.4 | 34.959 | 47 | 36.9589 | 42.4137 | 49.7 | 27 | 40.15 | 49.9 | 36.87 | 40.3 | 29.62 | 34.66 | 39.78 | [9 empty] | 36
| 1 | 3 | 27.3 | 30 | 37.1 | 33.4 | 31.3 | 32.7 | 31.7 | 34.746 | 50 | 37.4492 | 42.1866 | 50.5 | 27 | 40.26 | 49.9 | 35.76 | 42.2 | 30.32 | 35.35 | 41.21 | [9 empty] | 16
| 1 | 4 | 27.9 | 30.3 | 32.9 | 33.6 | 32.4 | 32.2 | 30.5 | 33.75 | 48 | 37.6432 | 41.8608 | 50.5 | 28 | 40.24 | 52.5 | 38.01 | 41.3 | 31.35 | 32.95 | 41.13 | [9 empty] | 35.1863
| 1 | 5 | 25.5 | 29.9 | 33.8 | 33.1 | 31.6 | 32 | 32.2 | 34.563 | 48 | 37.6839 | 42.0841 | 50 | 28 | 40.41 | 51.5 | 36.31 | 38.6 | 29.72 | 35.53 | 37.67 | [9 empty] | 77
| 1 | 6 | 26.1 | 31.4 | 32 | 34 | 32.5 | 31.9 | 31.9 | 32.445 | 50 | 36.9272 | 41.8182 | 50.4 | 28 | 40.35 | 50.5 | 32.85 | 39.5 | 31.08 | 33.24 | 40.47
| Mean |  | 26.9 | 30.5 | 33.8167 | 33.1833 | 32.2667 | 32.35 | 31.4 | 34.4477 | 49 | 37.1008 | 42.1237 | 50.2667 | 27.5 | 40.33 | 51.35 | 36.0567 | 40.1167 | 30.3783 | 34.6517 | 40.4833
| Median |  | 26.95 | 30.45 | 33.35 | 33.35 | 32.45 | 32.1 | 31.55 | 34.6545 | 49 | 37.204 | 42.1354 | 50.45 | 27.5 | 40.305 | 51 | 36.425 | 39.9 | 30.25 | 35.005 | 40.8
| Std Dev. |  | 1.006 | 0.55857 | 1.9631 | 0.78846 | 0.7174 | 0.54681 | 0.67528 | 1.2659 | 1.5492 | 0.65537 | 0.25184 | 0.33862 | 0.54772 | 0.14819 | 1.5668 | 1.7404 | 1.4275 | 0.70536 | 1.3026 | 1.6737
| Rel.Std.Dev. |  | 3.74% | 1.831% | 5.805% | 2.376% | 2.223% | 1.69% | 2.151% | 3.675% | 3.162% | 1.766% | 0.5979% | 0.6737% | 1.992% | 0.3674% | 3.051% | 4.827% | 3.558% | 2.322% | 3.759% | 4.134%
| PDM3 |  | -23.55% | -13.32% | -3.893% | -5.692% | -8.298% | -8.061% | -10.76% | -2.099% | 39.26% | 5.441% | 19.72% | 42.86% | -21.84% | 14.62% | 45.94% | 2.474% | 14.01% | -13.66% | -1.519% | 15.05%
| Z-Score (Absolute) |  | 1.25 | 0.66 | 0.12 | 0.22 | 0.37 | 0.36 | 0.52 | 0.02 | 2.37 | 0.42 | 1.24 | 2.58 | 1.16 | 0.95 | 2.76 | 0.25 | 0.91 | 0.68 | 0.02 | 0.97
| Analytical results for Co in OREAS 627 (Certified Value 7.1 ppm) | [30 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 6 | 6.3 | 6.4 | 7.7 | 8.5 | 6.2 | 6.7 | 7.1 | 6.8 | 7 | 6.8 | 7 | 6.1211 | 8.03 | 6.9976 | 9.6 | 7.4 | 4.8 | 8.15 | 8.02 | 7.6 | 8 | 7.3 | 7.2 | 7.2 | 8 | 7.7 |  |  | 1
| 1 | 2 | 6 | 6 | 6.3 | 7.7 | 7.8 | 6.2 | 6.5 | 7.3 | 6.8 | 7.1 | 6.8 | 7 | 6.2012 | 8.26 | 7.0467 | 9.4 | 7.3 | 5 | 9.36 | 7.71 | 7.6 | 8 | 7.3 | 7.6 | 6.9 | 8 | 6.8 |  |  | 37
| 1 | 3 | 6 | 6.1 | 6.9 | 7.6 | 8.7 | 6 | 6.4 | 7.2 | 6.9 | 7.2 | 6.5 | 7 | 6.4209 | 7.99 | 6.8518 | 9 | 7.4 | 5 | 9.8 | 7.78 | 7.6 | 8 | 7.3 | 7.5 | 7 | 9 | 6.9 |  |  | 16
| 1 | 4 | 6 | 5.9 | 6.7 | 7.8 | 7.1 | 6.1 | 6.4 | 7.3 | 6.7 | 7.1 | 6.7 | 7 | 6.4915 | 8.37 | 6.8518 | 9.3 | 7.4 | 4.9 | 9.28 | 7.82 | 7.4 | 8 | 7.1 | 7.6 | 7.1 | 9 | 7.3 |  |  | 7.0991
| 1 | 5 | 6 | 6 | 6.7 | 7.7 | 8.1 | 6 | 6.3 | 7.1 | 7.1 | 7.1 | 7 | 7 | 6.3743 | 8.02 | 7.0642 | 9.5 | 7.3 | 5.1 | 9.49 | 7.92 | 7.7 | 8 | 7 | 7.7 | 6.9 | 9 | 7.4 |  |  | 78
| 1 | 6 | 6 | 6 | 6.2 | 8 | 7.9 | 6.2 | 6.5 | 7.3 | 6.9 | 6.7 | 6.7 | 7 | 6.3045 | 7.83 | 7.031 | 9.2 | 7.4 | 4.9 | 9.75 | 7.76 | 7.3 | 8 | 6.8 | 7.7 | 7 | 9 | 7.9
| Mean |  | 6 | 6.05 | 6.5333 | 7.75 | 8.0167 | 6.1167 | 6.4667 | 7.2167 | 6.8667 | 7.0333 | 6.75 | 7 | 6.3189 | 8.0833 | 6.9739 | 9.3333 | 7.3667 | 4.95 | 9.305 | 7.835 | 7.5333 | 8 | 7.1333 | 7.55 | 7.0167 | 8.6667 | 7.3333
| Median |  | 6 | 6 | 6.55 | 7.7 | 8 | 6.15 | 6.45 | 7.25 | 6.85 | 7.1 | 6.75 | 7 | 6.3394 | 8.025 | 7.0143 | 9.35 | 7.4 | 4.95 | 9.425 | 7.8 | 7.6 | 8 | 7.2 | 7.6 | 7 | 9 | 7.35
| Std Dev. |  | 0 | 0.13784 | 0.27325 | 0.13784 | 0.56716 | 0.098319 | 0.13663 | 0.098319 | 0.13663 | 0.17512 | 0.16432 | 0 | 0.1389 | 0.19654 | 0.097059 | 0.21602 | 0.05164 | 0.10488 | 0.60249 | 0.11485 | 0.15055 | 0 | 0.20656 | 0.18708 | 0.1169 | 0.5164 | 0.43205
| Rel.Std.Dev. |  | 0% | 2.278% | 4.182% | 1.779% | 7.075% | 1.607% | 2.113% | 1.362% | 1.99% | 2.49% | 2.434% | 0% | 2.198% | 2.431% | 1.392% | 2.315% | 0.701% | 2.119% | 6.475% | 1.466% | 1.999% | 0% | 2.896% | 2.478% | 1.666% | 5.958% | 5.892%
| PDM3 |  | -15.48% | -14.78% | -7.969% | 9.169% | 12.93% | -13.84% | -8.908% | 1.657% | -3.274% | -0.9258% | -4.917% | -1.395% | -10.99% | 13.86% | -1.764% | 31.47% | 3.77% | -30.27% | 31.07% | 10.37% | 6.117% | 12.69% | 0.4829% | 6.352% | -1.161% | 22.08% | 3.3%
| Z-Score (Absolute) |  | NA | 1.22 | 0.67 | 0.69 | 0.99 | 1.14 | 0.75 | 0.09 | 0.3 | 0.11 | 0.43 | NA | 0.92 | 1.07 | 0.18 | 2.47 | 0.26 | 2.45 | 2.44 | 0.79 | 0.45 | NA | 0 | 0.47 | 0.13 | NA | 0.22
| Analytical results for Cr in OREAS 627 (Certified Value 15.9 ppm) | [30 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 2
| 1 | 1 | 17 | 13 | 14 | 14 | 16 | 14 | 15 | 15 | 16.1 | 16 | 15.7 | 16 | 15.0121 | 18.54 | 16.3004 | 18 | 14 | 14.79 | 17 | 15.98 | 17 | 31 | 16 | 16 | 16 | 18 | 17 |  |  | 1
| 1 | 2 | 15 | 12 | 14 | 15 | 15 | 15 | 14 | 16 | 17.3 | 16 | 15.9 | 14 | 16.4405 | 18.22 | 16.3362 | 17 | 16 | 14.83 | 17 | 15.06 | 17 | 30 | 16 | 15 | 15 | 18 | 17 |  |  | 15
| 1 | 3 | 17 | 12 | 14 | 15 | 17 | 15 | 14 | 16 | 17.6 | 15 | 15.3 | 16 | 15.5743 | 16.41 | 15.6348 | 17 | 15 | 14.94 | 19 | 15.43 | 17 | 30 | 17 | 17 | 16 | 18 | 17 |  |  | 16
| 1 | 4 | 17 | 12 | 14 | 15 | 15 | 16 | 15 | 16 | 16.6 | 15 | 15.8 | 15 | 16.5282 |  | 15.9024 | 17 | 15 | 14.18 | 17 | 15.12 | 17 | 32 | 16 | 16 | 16 | 19 | 17 |  |  | 15.9442
| 1 | 5 | 18 | 11 | 14 | 15 | 16 | 15 | 14 | 16 | 16.7 | 15 | 16.6 | 15 | 16.1133 | 19.5 | 16.3279 | 17 | 16 | 14.73 | 18 | 15.45 | 16 | 32 | 17 | 16 | 16 | 16 | 16 |  |  | 79
| 1 | 6 | 17 | 12 | 14 | 16 | 16 | 15 | 14 | 15 | 16.8 | 13 | 15.7 | 15 | 15.352 | 16.88 | 15.6911 | 18 | 15 | 14.39 | 18 | 14.87 | 16 | 30 | 17 | 16 | 16 | 17 | 16
| Mean |  | 16.8333 | 12 | 14 | 15 | 15.8333 | 15 | 14.3333 | 15.6667 | 16.85 | 15 | 15.8333 | 15.1667 | 15.8367 | 17.91 | 16.0321 | 17.3333 | 15.1667 | 14.6433 | 17.6667 | 15.3183 | 16.6667 | 30.8333 | 16.5 | 16 | 15.8333 | 17.6667 | 16.6667
| Median |  | 17 | 12 | 14 | 15 | 16 | 15 | 14 | 16 | 16.75 | 15 | 15.75 | 15 | 15.8438 | 18.22 | 16.1014 | 17 | 15 | 14.76 | 17.5 | 15.275 | 17 | 30.5 | 16.5 | 16 | 16 | 18 | 17
| Std Dev. |  | 0.98319 | 0.63246 | 0 | 0.63246 | 0.75277 | 0.63246 | 0.5164 | 0.5164 | 0.53198 | 1.0954 | 0.4274 | 0.75277 | 0.61691 | 1.2582 | 0.32954 | 0.5164 | 0.75277 | 0.29351 | 0.8165 | 0.39362 | 0.5164 | 0.98319 | 0.54772 | 0.63246 | 0.40825 | 1.0328 | 0.5164
| Rel.Std.Dev. |  | 5.841% | 5.27% | 0% | 4.216% | 4.754% | 4.216% | 3.603% | 3.296% | 3.157% | 7.303% | 2.699% | 4.963% | 3.895% | 7.025% | 2.055% | 2.979% | 4.963% | 2.004% | 4.622% | 2.57% | 3.098% | 3.189% | 3.32% | 3.953% | 2.578% | 5.846% | 3.098%
| PDM3 |  | 5.577% | -24.74% | -12.19% | -5.922% | -0.6951% | -5.922% | -10.1% | -1.74% | 5.681% | -5.922% | -0.6951% | -4.876% | -0.6738% | 12.33% | 0.5518% | 8.713% | -4.876% | -8.159% | 10.8% | -3.925% | 4.532% | 93.38% | 3.486% | 0.3503% | -0.6951% | 10.8% | 4.532%
| Z-Score (Absolute) |  | 0.81 | 3.1 | 1.48 | 0.67 | 0 | 0.67 | 1.21 | 0.13 | 0.82 | 0.67 | 0 | 0.54 | 0 | 1.68 | 0.16 | 1.21 | 0.54 | 0.96 | 1.48 | 0.42 | 0.67 | 12.14 | 0.54 | 0.13 | 0 | 1.48 | 0.67
| Analytical results for Cs in OREAS 627 (Certified Value 1.67 ppm) | [30 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 2
| 1 | 1 | 1.55 | 1.37 | 1.86 | 1.75 | 1.67 | 1.62 | 1.61 | 2.15 | 1.4609 | 1.6522 | 1.98 | 1.3 | 2.3 | 1.82 | 1.58 | 1.63 | 1.86 | 1.92 | [11 empty] | 1
| 1 | 2 | 1.46 | 1.41 | 1.65 | 1.78 | 1.6 | 1.68 | 1.63 | 2.03 | 1.5107 | 1.6308 | 1.95 | 1.4 | 2.21 | 1.87 | 1.67 | 1.69 | 1.76 | 1.75 | [11 empty] | 39
| 1 | 3 | 1.6 | 1.38 | 1.87 | 1.78 | 1.51 | 1.62 | 1.67 | 2 | 1.5618 | 1.5717 | 1.98 | 1.4 | 2.16 | 1.83 | 1.67 | 1.73 | 1.79 | 1.84 | [11 empty] | 16
| 1 | 4 | 1.47 | 1.42 | 1.68 | 1.77 | 1.56 | 1.64 | 1.63 | 1.91 | 1.5619 | 1.5322 | 1.93 | 1.4 | 2.2 | 1.8 | 1.59 | 1.73 | 1.69 | 1.82 | [11 empty] | 1.6743
| 1 | 5 | 1.44 | 1.39 | 1.73 | 1.82 | 1.52 | 1.66 | 1.69 | 1.94 | 1.537 | 1.5681 | 1.95 | 1.4 | 2.29 | 1.88 | 1.58 | 1.73 | 1.75 | 1.72 | [11 empty] | 80
| 1 | 6 | 1.47 | 1.47 | 1.63 | 1.82 | 1.56 | 1.58 | 1.71 | 1.88 | 1.5142 | 1.5156 | 1.99 | 1.4 | 2.26 | 1.75 | 1.55 | 1.78 | 1.87 | 1.81
| Mean |  | 1.4983 | 1.4067 | 1.7367 | 1.7867 | 1.57 | 1.6333 | 1.6567 | 1.985 | 1.5244 | 1.5784 | 1.9633 | 1.3833 | 2.2367 | 1.825 | 1.6067 | 1.715 | 1.7867 | 1.81
| Median |  | 1.47 | 1.4 | 1.705 | 1.78 | 1.56 | 1.63 | 1.65 | 1.97 | 1.5256 | 1.5699 | 1.965 | 1.4 | 2.235 | 1.825 | 1.585 | 1.73 | 1.775 | 1.815
| Std Dev. |  | 0.062423 | 0.036148 | 0.10501 | 0.028048 | 0.058652 | 0.035024 | 0.039328 | 0.098133 | 0.038161 | 0.053687 | 0.023381 | 0.040825 | 0.055377 | 0.047645 | 0.050859 | 0.050498 | 0.068896 | 0.070427
| Rel.Std.Dev. |  | 4.166% | 2.57% | 6.047% | 1.57% | 3.736% | 2.144% | 2.374% | 4.944% | 2.503% | 3.401% | 1.191% | 2.951% | 2.476% | 2.611% | 3.166% | 2.944% | 3.856% | 3.891%
| PDM3 |  | -10.51% | -15.98% | 3.726% | 6.712% | -6.229% | -2.446% | -1.052% | 18.56% | -8.951% | -5.725% | 17.26% | -17.38% | 33.59% | 9.002% | -4.039% | 2.432% | 6.712% | 8.106%
| Z-Score (Absolute) |  | 1.05 | 1.57 | 0.29 | 0.57 | 0.65 | 0.3 | 0.16 | 1.68 | 0.91 | 0.6 | 1.56 | 1.7 | 3.1 | 0.78 | 0.44 | 0.16 | 0.57 | 0.7
| Analytical results for Cu in OREAS 627 (Certified Value 0.481 wt.%) | [30 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 0.4871 | 0.431 | 0.47792 | 0.472 | 0.513 | 0.454 | 0.486 | 0.497 | 0.456 | 0.46035 | 0.488 | 0.4788 | 0.4555 | 0.44907 | 0.5176 | 0.468 | 0.50168 | 0.4646 | 0.5251 | 0.4849 | 0.52151 | 0.4899 | 0.4793 | 0.4737 | 0.48724 | 0.46963 | 0.476 | 0.45677 |  | 1
| 1 | 2 | 0.4688 | 0.443 | 0.4744 | 0.478 | 0.478 | 0.457 | 0.483 | 0.501 | 0.458 | 0.47069 | 0.488 | 0.4601 | 0.47266 | 0.47253 | 0.5109 | 0.463 | 0.49803 | 0.4626 | 0.5258 | 0.4875 | 0.52141 | 0.4711 | 0.4845 | 0.4783 | 0.47967 | 0.47354 | 0.4919 | 0.40817 |  | 25
| 1 | 3 | 0.4859 | 0.446 | 0.48288 | 0.479 | 0.526 | 0.464 | 0.482 | 0.492 | 0.47 | 0.46623 | 0.482 | 0.4733 | 0.46531 | 0.44824 | 0.5102 | 0.458 | 0.50694 | 0.4547 | 0.5166 | 0.4867 | 0.52078 | 0.4838 | 0.4708 | 0.4781 | 0.48159 | 0.48549 | 0.4946 | 0.41411 |  | 16
| 1 | 4 | 0.4896 | 0.405 | 0.48226 | 0.477 | 0.504 | 0.477 | 0.479 | 0.505 | 0.454 | 0.47285 | 0.482 | 0.4777 | 0.47334 | 0.4628 | 0.518 | 0.454 | 0.50494 | 0.4588 | 0.5314 | 0.4852 | 0.5328 | 0.4766 | 0.4844 | 0.479 | 0.48444 | 0.4891 | 0.4994 | 0.43851 |  | 0.4809
| 1 | 5 | 0.4897 | 0.428 | 0.47626 | 0.482 | 0.505 | 0.463 | 0.495 | 0.496 | 0.472 | 0.46436 | 0.479 | 0.4657 | 0.4726 | 0.47201 | 0.5161 | 0.464 | 0.50896 | 0.4596 | 0.528 | 0.4846 | 0.52366 | 0.4747 | 0.4745 | 0.4798 | 0.48282 | 0.48785 | 0.4953 | 0.45918 |  | 81
| 1 | 6 | 0.4916 | 0.446 | 0.47724 | 0.489 | 0.494 | 0.472 | 0.478 | 0.493 | 0.462 | 0.44606 | 0.485 | 0.4788 | 0.47202 | 0.45806 | 0.5125 | 0.458 | 0.50631 | 0.4619 | 0.5244 | 0.4866 | 0.52759 | 0.4837 | 0.4847 | 0.4795 | 0.48582 | 0.47364 | 0.4949 | 0.45745
| Mean |  | 0.48545 | 0.43317 | 0.47849 | 0.4795 | 0.50333 | 0.4645 | 0.48383 | 0.49733 | 0.462 | 0.46342 | 0.484 | 0.4724 | 0.46857 | 0.46045 | 0.51422 | 0.46083 | 0.50448 | 0.46037 | 0.52522 | 0.48592 | 0.52463 | 0.47997 | 0.4797 | 0.47807 | 0.4836 | 0.47987 | 0.49202 | 0.43903
| Median |  | 0.48835 | 0.437 | 0.47758 | 0.4785 | 0.5045 | 0.4635 | 0.4825 | 0.4965 | 0.46 | 0.4653 | 0.4835 | 0.4755 | 0.47231 | 0.46043 | 0.5143 | 0.4605 | 0.50562 | 0.46075 | 0.52545 | 0.4859 | 0.52259 | 0.48015 | 0.48185 | 0.47865 | 0.48363 | 0.47957 | 0.49475 | 0.44764
| Std Dev. |  | 0.0084054 | 0.015817 | 0.0033786 | 0.0056833 | 0.016367 | 0.008735 | 0.0061779 | 0.0049261 | 0.0074833 | 0.0096057 | 0.0036332 | 0.0078358 | 0.0070579 | 0.010669 | 0.0034464 | 0.0050761 | 0.0039781 | 0.003476 | 0.0049244 | 0.001172 | 0.0047192 | 0.0069999 | 0.0059421 | 0.0022385 | 0.002792 | 0.0085346 | 0.0082077 | 0.022951
| Rel.Std.Dev. |  | 1.731% | 3.651% | 0.7061% | 1.185% | 3.252% | 1.881% | 1.277% | 0.9905% | 1.62% | 2.073% | 0.7507% | 1.659% | 1.506% | 2.317% | 0.6702% | 1.102% | 0.7886% | 0.7551% | 0.9376% | 0.2412% | 0.8995% | 1.458% | 1.239% | 0.4682% | 0.5773% | 1.779% | 1.668% | 5.228%
| PDM3 |  | 0.9453% | -9.927% | -0.5013% | -0.292% | 4.664% | -3.411% | 0.6091% | 3.416% | -3.931% | -3.635% | 0.6438% | -1.768% | -2.565% | -4.253% | 6.927% | -4.174% | 4.902% | -4.271% | 9.214% | 1.042% | 9.092% | -0.1949% | -0.2504% | -0.59% | 0.5599% | -0.214% | 2.311% | -8.707%
| Z-Score (Absolute) |  | 0.28 | 2.28 | 0.06 | 0.01 | 1.15 | 0.75 | 0.2 | 0.86 | 0.87 | 0.8 | 0.21 | 0.36 | 0.55 | 0.95 | 1.69 | 0.93 | 1.21 | 0.95 | 2.23 | 0.3 | 2.2 | 0.01 | 0 | 0.08 | 0.19 | 0 | 0.6 | 2
| Analytical results for Dy in OREAS 627 (Indicative Value 1.59 ppm) | [30 empty] | Indicative
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  | 2
| 1 | 1 | 1.01: 1.8; 1.02: 1.816; 1.03: 1.371; 1.04: 1.4963 |  | 1
| 1 | 2 | 1.01: 1.7; 1.02: 1.754; 1.03: 1.387; 1.04: 1.479 |  | 7
| 1 | 3 | 1.01: 1.8; 1.02: 1.734; 1.03: 1.4311; 1.04: 1.6648 |  | 16
| 1 | 4 | 1.01: 1.8; 1.02: 1.729; 1.03: 1.4359; 1.04: 1.3729 |  | 1.5909
| 1 | 5 | 1.01: 1.7; 1.02: 1.735; 1.03: 1.4449; 1.04: 1.4202 |  | 13
| 1 | 6 | 1.01: 1.8; 1.02: 1.651; 1.03: 1.4256; 1.04: 1.4532
| Mean |  | 1.01: 1.7667; 1.02: 1.7365; 1.03: 1.4159; 1.04: 1.4811
| Median |  | 1.01: 1.8; 1.02: 1.7345; 1.03: 1.4283; 1.04: 1.4661
| Std Dev. |  | 1.01: 0.05164; 1.02: 0.052918; 1.03: 0.029728; 1.04: 0.10018
| Rel.Std.Dev. |  | 1.01: 2.923%; 1.02: 3.047%; 1.03: 2.1%; 1.04: 6.764%
| PDM3 |  | 1.01: 11.05%; 1.02: 9.155%; 1.03: -11%; 1.04: -6.901%
| Z-Score (Absolute) |  | 1.01: 0.75; 1.02: 0.6; 1.03: 0.91; 1.04: 0.6
| Analytical results for Er in OREAS 627 (Indicative Value 0.61 ppm) | [30 empty] | Indicative
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  | 2
| 1 | 1 | 1.01: 0.7; 1.02: 0.66; 1.03: 0.53709; 1.04: 0.60329 |  | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.628; 1.03: 0.51674; 1.04: 0.58443 |  | 8
| 1 | 3 | 1.01: 0.7; 1.02: 0.635; 1.03: 0.52798; 1.04: 0.65893 |  | 16
| 1 | 4 | 1.01: 0.7; 1.02: 0.643; 1.03: 0.54676; 1.04: 0.56305 |  | 0.61063
| 1 | 5 | 1.01: 0.6; 1.02: 0.648; 1.03: 0.51589; 1.04: 0.53962 |  | 14
| 1 | 6 | 1.01: 0.7; 1.02: 0.581; 1.03: 0.52863; 1.04: 0.57596
| Mean |  | 1.01: 0.68333; 1.02: 0.6325; 1.03: 0.52885; 1.04: 0.58754
| Median |  | 1.01: 0.7; 1.02: 0.639; 1.03: 0.52831; 1.04: 0.58019
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.027515; 1.03: 0.011866; 1.04: 0.04095
| Rel.Std.Dev. |  | 1.01: 5.974%; 1.02: 4.35%; 1.03: 2.244%; 1.04: 6.97%
| PDM3 |  | 1.01: 11.91%; 1.02: 3.581%; 1.03: -13.39%; 1.04: -3.781%
| Z-Score (Absolute) |  | 1.01: 1.03; 1.02: 0.32; 1.03: 1.14; 1.04: 0.32
| Analytical results for Eu in OREAS 627 (Indicative Value 0.67 ppm) | [30 empty] | Indicative
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  | 2
| 1 | 1 | 1.01: 0.7; 1.02: 0.779; 1.03: 0.66206; 1.04: 0.55951 |  | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.722; 1.03: 0.70042; 1.04: 0.5991 |  | 9
| 1 | 3 | 1.01: 0.7; 1.02: 0.755; 1.03: 0.67923; 1.04: 0.65322 |  | 16
| 1 | 4 | 1.01: 0.7; 1.02: 0.723; 1.03: 0.71225; 1.04: 0.5214 |  | 0.66558
| 1 | 5 | 1.01: 0.6; 1.02: 0.725; 1.03: 0.67528; 1.04: 0.50177 |  | 15
| 1 | 6 | 1.01: 0.7; 1.02: 0.673; 1.03: 0.69334; 1.04: 0.53938
| Mean |  | 1.01: 0.68333; 1.02: 0.7295; 1.03: 0.6871; 1.04: 0.5624
| Median |  | 1.01: 0.7; 1.02: 0.724; 1.03: 0.68629; 1.04: 0.54945
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.035809; 1.03: 0.018318; 1.04: 0.05566
| Rel.Std.Dev. |  | 1.01: 5.974%; 1.02: 4.909%; 1.03: 2.666%; 1.04: 9.897%
| PDM3 |  | 1.01: 2.667%; 1.02: 9.603%; 1.03: 3.233%; 1.04: -15.5%
| Z-Score (Absolute) |  | 1.01: 0.05; 1.02: 1.29; 1.03: 0.05; 1.04: 3.59
| Analytical results for Fe in OREAS 627 (Certified Value 6.34 wt.%) | [30 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 3
| 1 | 1 | 6.51 | 5.58 | 5.69 | 6.09 | 6.66 | 6.05 | 6.15 | 6.25 | 6.03 | 6.77 | 5.94 | 6.66 | 5.9905 | 8.13 | 6.6764 | 6.7 | 6.65 | 6.05 | 6.58 | 6.2256 | 6.51 | 6.49 | 6.63 | 6.35 | 6.38 | 6.4 | 6.22 | 6.38 |  | 1
| 1 | 2 | 6.56 | 5.26 | 5.66 | 6.3 | 6.23 | 6.14 | 6.07 | 6.35 | 6.13 | 6.59 | 5.88 | 6.43 | 6.2832 | 7.987 | 6.7045 | 6.68 | 6.64 | 6.03 | 6.5945 | 6.2833 | 6.55 | 6.28 | 6.55 | 6.31 | 6.27 | 6.29 | 6.25 | 5.81 |  | 26
| 1 | 3 | 6.58 | 5.49 | 5.8 | 6.16 | 6.93 | 6.18 | 5.95 | 6.29 | 6.45 | 6.39 | 5.79 | 6.59 | 6.1533 | 8.43 | 6.8202 | 6.73 | 6.52 | 6.04 | 6.6047 | 6.1752 | 6.4 | 6.42 | 6.43 | 6.32 | 6.39 | 6.58 | 6.33 | 5.82 |  | 16
| 1 | 4 | 6.52 | 5.37 | 5.8 | 6.17 | 6.27 | 6.28 | 6.01 | 6.37 | 6.14 | 6.36 | 5.84 | 6.68 | 6.235 | 9.602 | 6.7902 | 6.61 | 6.7 | 6.08 | 6.6035 | 6.3283 | 6.39 | 6.34 | 6.64 | 6.31 | 6.3 | 6.74 | 6.28 | 6.15 |  | 6.3378
| 1 | 5 | 6.55 | 5.24 | 5.76 | 6.19 | 6.55 | 6.1 | 6.04 | 6.28 | 6.24 | 6.33 | 6.05 | 6.51 | 6.282 | 9.312 | 6.7502 | 6.66 | 6.57 | 6.13 | 6.6025 | 6.2458 | 6.61 | 6.31 | 6.51 | 6.33 | 6.33 | 6.47 | 6.26 | 6.14 |  | 82
| 1 | 6 | 6.42 | 5.31 | 5.77 | 6.25 | 6.38 | 6.25 | 6 | 6.25 | 6.04 | 5.65 | 5.96 | 6.66 | 6.1836 | 7.949 | 6.7586 | 6.69 | 6.6 | 6.07 | 6.5987 | 6.1655 | 6.5 | 6.38 | 6.6 | 6.48 | 6.37 | 6.55 | 6.28 | 6.53
| Mean |  | 6.5233 | 5.375 | 5.7467 | 6.1933 | 6.5033 | 6.1667 | 6.0367 | 6.2983 | 6.1717 | 6.3483 | 5.91 | 6.5883 | 6.1879 | 8.5683 | 6.75 | 6.6783 | 6.6133 | 6.0667 | 6.5973 | 6.2373 | 6.4933 | 6.37 | 6.56 | 6.35 | 6.34 | 6.505 | 6.27 | 6.1383
| Median |  | 6.535 | 5.34 | 5.765 | 6.18 | 6.465 | 6.16 | 6.025 | 6.285 | 6.135 | 6.375 | 5.91 | 6.625 | 6.2093 | 8.28 | 6.7544 | 6.685 | 6.62 | 6.06 | 6.6006 | 6.2357 | 6.505 | 6.36 | 6.575 | 6.325 | 6.35 | 6.51 | 6.27 | 6.145
| Std Dev. |  | 0.056804 | 0.13487 | 0.058538 | 0.073394 | 0.26576 | 0.088015 | 0.068605 | 0.050761 | 0.15639 | 0.38107 | 0.092952 | 0.099883 | 0.10982 | 0.71474 | 0.053124 | 0.040702 | 0.06377 | 0.036148 | 0.0092664 | 0.062638 | 0.085479 | 0.076942 | 0.080498 | 0.065422 | 0.048166 | 0.15579 | 0.036878 | 0.2902
| Rel.Std.Dev. |  | 0.8708% | 2.509% | 1.019% | 1.185% | 4.086% | 1.427% | 1.136% | 0.8059% | 2.534% | 6.003% | 1.573% | 1.516% | 1.775% | 8.342% | 0.787% | 0.6095% | 0.9643% | 0.5958% | 0.1405% | 1.004% | 1.316% | 1.208% | 1.227% | 1.03% | 0.7597% | 2.395% | 0.5882% | 4.728%
| PDM3 |  | 2.927% | -15.19% | -9.328% | -2.28% | 2.611% | -2.701% | -4.752% | -0.6234% | -2.622% | 0.1655% | -6.751% | 3.952% | -2.365% | 35.19% | 6.504% | 5.372% | 4.347% | -4.279% | 4.094% | -1.587% | 2.453% | 0.5074% | 3.505% | 0.1918% | 0.03405% | 2.637% | -1.07% | -3.148%
| Z-Score (Absolute) |  | 0.68 | 3.66 | 2.26 | 0.57 | 0.6 | 0.67 | 1.16 | 0.17 | 0.65 | 0.02 | 1.64 | 0.92 | 0.59 | 8.41 | 1.53 | 1.26 | 1.02 | 1.05 | 0.96 | 0.4 | 0.56 | 0.1 | 0.82 | 0.02 | 0.02 | 0.61 | 0.28 | 0.78
| Analytical results for Ga in OREAS 627 (Certified Value 3.97 ppm) | [30 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g |  |  |  |  | 2
| 1 | 1 | 5 | 4.2 | 9 | 3.67 | 4.31 | 3.29 | 3.43 | 4.41 | 3.53 | 3.95 | 3.5 | 3.5765 | 5.29 | 4.1567 | 124.9 | 4.3 | 3 | < 5 | 6.72 | 4.5 | < 5 | 3.9 | 3.6 | 4.09 | 6.33 |  |  |  |  | 1
| 1 | 2 | 4 | 4.05 | 10 | 3.73 | 3.95 | 3.45 | 3.47 | 4.52 | 3.56 | 4.09 | 3.5 | 3.6618 | 6.132 | 4.1307 | 127 | 4.3 | 3 | < 5 | 6.18 | 4.7 | < 5 | 4.1 | 3.7 | 4.1 | 5.84 |  |  |  |  | 44
| 1 | 3 | 4 | 4.3 | 10 | 3.65 | 4.38 | 3.44 | 3.29 | 4.36 | 3.56 | 3.86 | 3.3 | 3.6592 | 5.45 | 4.156 | 126.3 | 4.5 | 3 | < 5 | 6.18 | 4.6 | < 5 | 4.3 | 3.6 | 3.95 | 6.36 |  |  |  |  | 16
| 1 | 4 | 5 | 4.01 | 10 | 3.73 | 3.82 | 3.27 | 3.39 | 4.45 | 3.58 | 3.79 | 3.4 | 3.7086 | 5.3 | 4.1301 | 130.2 | 4.3 | 3 | < 5 | 6.27 | 4.4 | < 5 | 4.1 | 3.8 | 4.04 | 6.3 |  |  |  |  | 3.9743
| 1 | 5 | 5 | 3.94 | 11 | 3.73 | 4.08 | 3.33 | 3.29 | 4.43 | 3.64 | 3.73 | 3.5 | 3.6454 | 4.53 | 4.1653 | 128.8 | 4.5 | 3 | < 5 | 6.51 | 4.8 | < 5 | 4 | 3.8 | 4.13 | 3.4 |  |  |  |  | 83
| 1 | 6 | 5 | 4.18 | 10 | 3.91 | 3.95 | 3.34 | 3.36 | 4.37 | 3.65 | 3.63 | 3.4 | 3.5884 | 5.99 | 4.1398 | 125.7 | 4.4 | 3 | < 5 | 6.32 | 4.4 | < 5 | 3.9 | 3.7 | 3.85 | 6.21
| Mean |  | 4.6667 | 4.1133 | 10 | 3.7367 | 4.0817 | 3.3533 | 3.3717 | 4.4233 | 3.5867 | 3.8417 | 3.4333 | 3.64 | 5.4487 | 4.1465 | 127.15 | 4.3833 | 3 |  | 6.3633 | 4.5667 |  | 4.05 | 3.7 | 4.0267 | 5.74
| Median |  | 5 | 4.115 | 10 | 3.73 | 4.015 | 3.335 | 3.375 | 4.42 | 3.57 | 3.825 | 3.45 | 3.6523 | 5.375 | 4.1479 | 126.65 | 4.35 | 3 |  | 6.295 | 4.55 |  | 4.05 | 3.7 | 4.065 | 6.255
| Std Dev. |  | 0.5164 | 0.1353 | 0.63246 | 0.091797 | 0.22104 | 0.075542 | 0.073326 | 0.058538 | 0.048028 | 0.16351 | 0.08165 | 0.049524 | 0.57467 | 0.014921 | 1.9967 | 0.098319 | 0 |  | 0.21285 | 0.1633 |  | 0.15166 | 0.089443 | 0.10708 | 1.1621
| Rel.Std.Dev. |  | 11.07% | 3.289% | 6.325% | 2.457% | 5.415% | 2.253% | 2.175% | 1.323% | 1.339% | 4.256% | 2.378% | 1.361% | 10.55% | 0.3599% | 1.57% | 2.243% | 0% |  | 3.345% | 3.576% |  | 3.745% | 2.417% | 2.659% | 20.25%
| PDM3 |  | 17.42% | 3.499% | 151.6% | -5.979% | 2.702% | -15.62% | -15.16% | 11.3% | -9.753% | -3.337% | -13.61% | -8.411% | 37.1% | 4.332% | 3099% | 10.29% | -24.51% |  | 60.11% | 14.91% |  | 1.905% | -6.901% | 1.318% | 44.43%
| Z-Score (Absolute) |  | NA | 0.12 | NA | 0.48 | 0.07 | 1.09 | 1.06 | 0.61 | 0.72 | 0.31 | 0.96 | 0.63 | 2.24 | 0.17 | 195.33 | 0.55 | NA | 2.44 | 3.69 | 0.84 | 2.44 | 0.02 | 0.54 | 0.02 | 2.7
Results from laboratories 1.01, 1.03 and 1.18 were removed due to their 1 ppm reading resolution.
| Analytical results for Gd in OREAS 627 (Indicative Value 2.64 ppm) | [30 empty] | Indicative
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  |  |  |  | 2
| 1 | 1 | 1.01: 2.5; 1.02: 2.79; 1.03: 2.4045; 1.04: 3.0656 |  |  |  |  | 1
| 1 | 2 | 1.01: 2.3; 1.02: 2.719; 1.03: 2.4648; 1.04: 3.0593 |  |  |  |  | 10
| 1 | 3 | 1.01: 2.4; 1.02: 2.735; 1.03: 2.5959; 1.04: 3.3077 |  |  |  |  | 16
| 1 | 4 | 1.01: 2.5; 1.02: 2.537; 1.03: 2.6083; 1.04: 2.9744 |  |  |  |  | 2.6427
| 1 | 5 | 1.01: 2.3; 1.02: 2.697; 1.03: 2.51; 1.04: 2.9931 |  |  |  |  | 16
| 1 | 6 | 1.01: 2.4; 1.02: 2.431; 1.03: 2.4746; 1.04: 2.9548
| Mean |  | 1.01: 2.4; 1.02: 2.6515; 1.03: 2.5097; 1.04: 3.0591
| Median |  | 1.01: 2.4; 1.02: 2.708; 1.03: 2.4923; 1.04: 3.0262
| Std Dev. |  | 1.01: 0.089443; 1.02: 0.13749; 1.03: 0.079333; 1.04: 0.12982
| Rel.Std.Dev. |  | 1.01: 3.727%; 1.02: 5.185%; 1.03: 3.161%; 1.04: 4.244%
| PDM3 |  | 1.01: -9.182%; 1.02: 0.3346%; 1.03: -5.031%; 1.04: 15.76%
| Z-Score (Absolute) |  | 1.01: 0.97; 1.02: 0.38; 1.03: 0.38; 1.04: 2.57
| Analytical results for Ge in OREAS 627 (Certified Value 0.095 ppm) | [30 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.04; 1.03: 1.05; 1.04: 1.06; 1.05: 1.07; 1.06: 1.08; 1.07: 1.09; 1.08: 1.13; 1.09: 1.18; 1.10: 1.23; 1.11: 1.25; 1.13: 1.26; 1.14: 1.27; 1.15: 1.30 |  |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.09: AR*OES/MS; 1.10: AR*MS; 1.11: AR*OES/MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS; 1.15: AR*MS |  |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 0.2g; 1.09: 0.25g; 1.10: 0.25g; 1.11: 0.25g; 1.13: 0.25g; 1.14: 0.5g; 1.15: 0.2g |  |  |  |  | 3
| 1 | 1 | 1.01: < 0.1; 1.02: 0.09; 1.03: 0.11; 1.04: 0.05; 1.05: 0.07; 1.06: 0.13; 1.07: 0.11; 1.08: < 0.2; 1.09: 0.5; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.07; 1.15: 0.23 |  |  |  |  | 1
| 1 | 2 | 1.01: < 0.1; 1.02: 0.09; 1.03: 0.1; 1.04: 0.06; 1.05: 0.08; 1.06: 0.17; 1.07: 0.12; 1.08: < 0.2; 1.09: 0.5; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.08; 1.15: 0.2 |  |  |  |  | 46
| 1 | 3 | 1.01: < 0.1; 1.02: 0.1; 1.03: 0.12; 1.04: 0.06; 1.05: 0.07; 1.06: 0.13; 1.07: 0.11; 1.08: < 0.2; 1.09: 0.6; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.07; 1.15: 0.2 |  |  |  |  | 16
| 1 | 4 | 1.01: < 0.1; 1.02: 0.09; 1.03: 0.11; 1.04: < 0.05; 1.05: 0.07; 1.06: 0.13; 1.07: 0.12; 1.08: < 0.2; 1.09: 0.5; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.07; 1.15: 0.22 |  |  |  |  | 0.095381
| 1 | 5 | 1.01: < 0.1; 1.02: 0.1; 1.03: 0.12; 1.04: 0.05; 1.05: 0.07; 1.06: 0.15; 1.07: 0.12; 1.08: < 0.2; 1.09: 0.6; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.08; 1.15: 0.2 |  |  |  |  | 84
| 1 | 6 | 1.01: < 0.1; 1.02: 0.1; 1.03: 0.1; 1.04: 0.06; 1.05: 0.08; 1.06: 0.14; 1.07: 0.11; 1.08: < 0.2; 1.09: 0.5; 1.10: < 0.1; 1.11: < 0.1; 1.13: 0.1; 1.14: 0.09; 1.15: 0.17
| Mean |  | 1.02: 0.095; 1.03: 0.11; 1.04: 0.056; 1.05: 0.073333; 1.06: 0.14167; 1.07: 0.115; 1.09: 0.53333; 1.13: 0.1; 1.14: 0.076667; 1.15: 0.20333
| Median |  | 1.02: 0.095; 1.03: 0.11; 1.04: 0.06; 1.05: 0.07; 1.06: 0.135; 1.07: 0.115; 1.09: 0.5; 1.13: 0.1; 1.14: 0.075; 1.15: 0.2
| Std Dev. |  | 1.02: 0.0054772; 1.03: 0.0089443; 1.04: 0.0054772; 1.05: 0.005164; 1.06: 0.016021; 1.07: 0.0054772; 1.09: 0.05164; 1.13: 1.5202e-17; 1.14: 0.008165; 1.15: 0.020656
| Rel.Std.Dev. |  | 1.02: 5.766%; 1.03: 8.131%; 1.04: 9.781%; 1.05: 7.042%; 1.06: 11.31%; 1.07: 4.763%; 1.09: 9.682%; 1.13: 1.52e-14%; 1.14: 10.65%; 1.15: 10.16%
| PDM3 |  | 1.02: -0.3994%; 1.03: 15.33%; 1.04: -41.29%; 1.05: -23.12%; 1.06: 48.53%; 1.07: 20.57%; 1.09: 459.2%; 1.13: 4.843%; 1.14: -19.62%; 1.15: 113.2%
| Z-Score (Absolute) |  | 1.01: 0.75; 1.02: 0.19; 1.03: 0.51; 1.04: 0.74; 1.05: 0.26; 1.06: 1.17; 1.07: 0.61; 1.08: 0.3; 1.09: NA; 1.10: 0.75; 1.11: 0.75; 1.13: NA; 1.14: 0.19; 1.15: 2.47
Results from laboratories 1.18 and 1.26 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Hf in OREAS 627 (Certified Value 1.07 ppm) | [30 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 2
| 1 | 1 | 1.4 | 0.91 | 1.14 | 0.94 | 1 | 1.07 | 0.73 | 1.2 | 0.87456 | 1.1137 | 1.28 | 1.1 | 1.44 | 1.24 | 1.01 | 0.87 | 0.95 | 0.94 | [11 empty] | 1
| 1 | 2 | 1.4 | 0.95 | 1.02 | 0.94 | 1.06 | 0.99 | 0.74 | 1.2 | 0.86143 | 1.1325 | 1.28 | 1.1 | 1.41 | 1.09 | 1.08 | 0.82 | 1.01 | 0.97 | [11 empty] | 16
| 1 | 3 | 1.5 | 0.92 | 1.15 | 0.96 | 0.97 | 1.03 | 0.75 | 1.19 | 0.88396 | 1.1256 | 1.3 | 1.1 | 1.36 | 1.18 | 1.12 | 0.85 | 0.96 | 1.14 | [11 empty] | 16
| 1 | 4 | 1.4 | 0.95 | 1 | 0.96 | 1.02 | 1.03 | 0.72 | 1.27 | 0.86551 | 1.1208 | 1.29 | 1.1 | 1.39 | 1.12 | 1.1 | 0.85 | 0.98 | 1.03 | [11 empty] | 1.0681
| 1 | 5 | 1.3 | 0.95 | 1.12 | 0.92 | 0.98 | 1.01 | 0.74 | 1.25 | 0.87132 | 1.1091 | 1.29 | 1.1 | 1.38 | 1.12 | 1.13 | 0.8 | 0.96 | 1.36 | [11 empty] | 85
| 1 | 6 | 1.4 | 0.94 | 1.11 | 0.98 | 0.99 | 1.03 | 0.71 | 1.12 | 0.86746 | 1.1245 | 1.29 | 1.1 | 1.37 | 1.2 | 1.07 | 0.85 | 0.98 | 0.88
| Mean |  | 1.4 | 0.93667 | 1.09 | 0.95 | 1.0033 | 1.0267 | 0.73167 | 1.205 | 0.87071 | 1.121 | 1.2883 | 1.1 | 1.3917 | 1.1583 | 1.085 | 0.84 | 0.97333 | 1.0533
| Median |  | 1.4 | 0.945 | 1.115 | 0.95 | 0.995 | 1.03 | 0.735 | 1.2 | 0.86939 | 1.1227 | 1.29 | 1.1 | 1.385 | 1.15 | 1.09 | 0.85 | 0.97 | 1
| Std Dev. |  | 0.063246 | 0.017512 | 0.063875 | 0.020976 | 0.03266 | 0.026583 | 0.01472 | 0.05244 | 0.0079298 | 0.0084852 | 0.0075277 | 2.4324e-16 | 0.029269 | 0.057417 | 0.043243 | 0.025298 | 0.021602 | 0.17432
| Rel.Std.Dev. |  | 4.518% | 1.87% | 5.86% | 2.208% | 3.255% | 2.589% | 2.012% | 4.352% | 0.9107% | 0.7569% | 0.5843% | 2.211e-14% | 2.103% | 4.957% | 3.986% | 3.012% | 2.219% | 16.55%
| PDM3 |  | 31.08% | -12.3% | 2.054% | -11.05% | -6.06% | -3.876% | -31.5% | 12.82% | -18.48% | 4.961% | 20.62% | 2.99% | 30.3% | 8.452% | 1.586% | -21.35% | -8.869% | -1.379%
| Z-Score (Absolute) |  | 2.08 | 0.83 | 0.13 | 0.75 | 0.41 | 0.27 | 2.12 | 0.85 | 1.24 | 0.33 | 1.37 | 0.19 | 2.02 | 0.56 | 0.1 | 1.44 | 0.6 | 0.1
| Analytical results for Hg in OREAS 627 (Certified Value 3.42 ppm) | [30 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.14 | 1.15 | 1.16 | 1.18 | 1.20 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | [5 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 0.5g | 01g | 01g | 0.25g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [5 empty] | 3
| 1 | 1 | < 5 | 3.26 | 3.29 | 3.43 | 3.34 | 3.55 | 3.43 | 3.62 | 3.34 | 3.9 | 3.432 | 3.6162 | 3.38 | 2.9473 | < 5 | 1.386 | 3.04 | 3.25 | 4 | 3.5 | 3.97 | 3.226 | 4 | 3 | [5 empty] | 1
| 1 | 2 | < 5 | 3.14 | 3.28 | 3.4 | 3.16 | 3.6 | 3.52 | 3.69 | 3.39 | 3.8 | 3.394 | 3.8724 | 2.8 | 3.2248 | < 5 | 1.336 | 3.05 | 3.04 | 4 | 3.7 | 3.87 | 3.109 | 5 | 3 | [5 empty] | 4
| 1 | 3 | < 5 | 3.16 | 3.54 | 3.29 | 3.54 | 3.35 | 3.33 | 3.56 | 3.45 | 3.8 | 3.259 | 3.6843 | 4.01 | 3.266 | < 5 | 1.35 | 3.02 | 3.34 | 4 | 3.7 | 4.13 | 3.183 | 5 | 3 | [5 empty] | 16
| 1 | 4 | < 5 | 3.16 | 3.63 | 3.46 | 3.3 | 3.39 | 3.36 | 3.54 | 3.42 | 3.8 | 3.32 | 3.7946 | 3.71 | 3.2887 | < 5 | 1.386 | 3.07 | 3.14 | 4 | 3.6 | 4.08 | 3.239 | 4 | 3 | [5 empty] | 3.4225
| 1 | 5 | < 5 | 2.91 | 3.51 | 3.46 | 3.26 | 3.44 | 3.32 | 3.52 | 3.43 | 3.8 | 3.472 | 3.7229 | 3.38 | 3.3335 | < 5 | 1.361 | 3.1 | 3.18 | 4 | 3.6 | 3.9 | 3.064 | 1 | 3 | [5 empty] | 86
| 1 | 6 | < 5 | 3.1 | 3.51 | 3.36 | 3.24 | 3.4 | 3.42 | 3.57 | 3.34 | 3.6 | 3.251 | 3.6869 | 3.3 | 2.9518 | < 5 | 1.381 | 3 | 3.17 | 3 | 3.5 | 4.01 | 3.088 | 4 | 3
| Mean |  |  | 3.1217 | 3.46 | 3.4 | 3.3067 | 3.455 | 3.3967 | 3.5833 | 3.395 | 3.7833 | 3.3547 | 3.7295 | 3.43 | 3.1687 |  | 1.3667 | 3.0467 | 3.1867 | 3.8333 | 3.6 | 3.9933 | 3.1515 | 3.8333 | 3
| Median |  |  | 3.15 | 3.51 | 3.415 | 3.28 | 3.42 | 3.39 | 3.565 | 3.405 | 3.8 | 3.357 | 3.7049 | 3.38 | 3.2454 |  | 1.371 | 3.045 | 3.175 | 4 | 3.6 | 3.99 | 3.146 | 4 | 3
| Std Dev. |  |  | 0.11635 | 0.14255 | 0.06603 | 0.12941 | 0.098539 | 0.075542 | 0.062183 | 0.046797 | 0.098319 | 0.092082 | 0.090978 | 0.40831 | 0.17335 |  | 0.020992 | 0.03559 | 0.10152 | 0.40825 | 0.089443 | 0.10093 | 0.074422 | 1.472 | 0
| Rel.Std.Dev. |  |  | 3.727% | 4.12% | 1.942% | 3.914% | 2.852% | 2.224% | 1.735% | 1.378% | 2.599% | 2.745% | 2.439% | 11.9% | 5.471% |  | 1.536% | 1.168% | 3.186% | 10.65% | 2.485% | 2.527% | 2.361% | 38.4% | 0%
| PDM3 |  |  | -8.79% | 1.096% | -0.6575% | -3.385% | 0.9495% | -0.7549% | 4.699% | -0.8036% | 10.54% | -1.982% | 8.971% | 0.2191% | -7.416% |  | -60.07% | -10.98% | -6.891% | 12% | 5.186% | 16.68% | -7.918% | 12% | -12.34%
| Z-Score (Absolute) |  | 2.9 | 0.88 | 0.21 | 0.02 | 0.29 | 0.19 | 0.01 | 0.61 | 0 | 1.26 | 0.13 | 1.08 | 0.11 | 0.73 | 2.9 | 6.57 | 1.13 | 0.67 | NA | 0.66 | 1.94 | 0.79 | NA | NA
Results from laboratories 1.24, 1.28 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Ho in OREAS 627 (Indicative Value 0.26 ppm) | [30 empty] | Indicative
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | [5 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | [5 empty] | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | [5 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | [5 empty] | 2
| 1 | 1 | 1.01: 0.3; 1.02: 0.283; 1.03: 0.22879; 1.04: 0.23239 | [5 empty] | 1
| 1 | 2 | 1.01: 0.3; 1.02: 0.276; 1.03: 0.23156; 1.04: 0.22406 | [5 empty] | 6
| 1 | 3 | 1.01: 0.3; 1.02: 0.261; 1.03: 0.23233; 1.04: 0.25908 | [5 empty] | 16
| 1 | 4 | 1.01: 0.3; 1.02: 0.267; 1.03: 0.23361; 1.04: 0.21481 | [5 empty] | 0.2572
| 1 | 5 | 1.01: 0.3; 1.02: 0.273; 1.03: 0.23156; 1.04: 0.20801 | [5 empty] | 12
| 1 | 6 | 1.01: 0.3; 1.02: 0.262; 1.03: 0.22765; 1.04: 0.22694
| Mean |  | 1.01: 0.3; 1.02: 0.27033; 1.03: 0.23092; 1.04: 0.22755
| Median |  | 1.01: 0.3; 1.02: 0.27; 1.03: 0.23156; 1.04: 0.2255
| Std Dev. |  | 1.01: 0; 1.02: 0.0085713; 1.03: 0.0022484; 1.04: 0.017744
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 3.171%; 1.03: 0.9737%; 1.04: 7.798%
| PDM3 |  | 1.01: 16.64%; 1.02: 5.106%; 1.03: -10.22%; 1.04: -11.53%
| Z-Score (Absolute) |  | 1.01: 1.56; 1.02: 0.62; 1.03: 0.62; 1.04: 0.73
| Analytical results for In in OREAS 627 (Certified Value 8.8 ppm) | [30 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.14 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [10 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.5g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [10 empty] | 3
| 1 | 1 | 7 | 7.97 | 9.91 | 9.1 | 9.67 | 9.3 | 9.27 | 8.52 | 8.94 | 15.571 | 7.7093 | 8.62 | 8.09 | 9.4 | 9.17 | 8.09 | 8.03 | 9.333 | 10.03 | [10 empty] | 1
| 1 | 2 | 5 | 7.85 | 9.94 | 8.59 | 10.1 | 9.38 | 9.43 | 8.54 | 8.81 | 14.752 | 7.7116 | 8.71 | 7.9 | 9.3 | 8.79 | 8.51 | 8.04 | 9.001 | 8.513 | [10 empty] | 33
| 1 | 3 | 5 | 7.88 | 9.83 | 9.37 | 9.68 | 9.05 | 9.17 | 8.8 | 8.58 | 18.018 | 7.7332 | 8.72 | 8.13 | 9.1 | 8.69 | 8.47 | 8.26 | 9.429 | 9.268 | [10 empty] | 16
| 1 | 4 | 8 | 7.72 | 9.86 | 9.32 | 9.43 | 9.16 | 9.36 | 8.61 | 8.49 | 15.5 | 7.8045 | 8.85 | 8 | 8.9 | 8.71 | 8.38 | 8.21 | 8.941 | 9.327 | [10 empty] | 8.8014
| 1 | 5 | 7 | 7.63 | 9.84 | 8.93 | 9.54 | 8.93 | 9.17 | 8.85 | 8.75 | 12.977 | 7.7976 | 8.58 | 8.16 | 9.4 | 8.89 | 8.4 | 8.23 | 8.915 | 9.128 | [10 empty] | 87
| 1 | 6 | 5 | 7.68 | 9.87 | 8.68 | 9.62 | 9.25 | 9.26 | 8.71 | 8.4 | 24.756 | 7.7474 | 8.69 | 8.09 | 9.1 | 8.93 | 8.46 | 8.18 | 9.058 | 9.531
| Mean |  | 6.1667 | 7.7883 | 9.875 | 8.9983 | 9.6733 | 9.1783 | 9.2767 | 8.6717 | 8.6617 | 16.929 | 7.7506 | 8.695 | 8.0617 | 9.2 | 8.8633 | 8.385 | 8.1583 | 9.1128 | 9.2995
| Median |  | 6 | 7.785 | 9.865 | 9.015 | 9.645 | 9.205 | 9.265 | 8.66 | 8.665 | 15.5355 | 7.7403 | 8.7 | 8.09 | 9.2 | 8.84 | 8.43 | 8.195 | 9.0295 | 9.2975
| Std Dev. |  | 1.3292 | 0.13167 | 0.042308 | 0.32394 | 0.22888 | 0.16678 | 0.10347 | 0.13703 | 0.20586 | 4.1639 | 0.041594 | 0.093541 | 0.095795 | 0.2 | 0.17784 | 0.15215 | 0.099079 | 0.21569 | 0.49739
| Rel.Std.Dev. |  | 21.55% | 1.691% | 0.4284% | 3.6% | 2.366% | 1.817% | 1.115% | 1.58% | 2.377% | 24.6% | 0.5367% | 1.076% | 1.188% | 2.174% | 2.006% | 1.815% | 1.214% | 2.367% | 5.349%
| PDM3 |  | -29.94% | -11.51% | 12.2% | 2.238% | 9.907% | 4.283% | 5.4% | -1.474% | -1.587% | 92.35% | -11.94% | -1.209% | -8.404% | 4.529% | 0.704% | -4.731% | -7.306% | 3.539% | 5.66%
| Z-Score (Absolute) |  | NA | 2.16 | 1.78 | 0.13 | 1.4 | 0.47 | 0.65 | 0.49 | 0.51 | 15.1 | 2.23 | 0.45 | 1.64 | 0.51 | 0.13 | 1.03 | 1.46 | 0.34 | 0.7
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for K in OREAS 627 (Certified Value 0.318 wt.%) | [30 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.32 | 0.29 | 0.28 | 0.27 | 0.36 | 0.29 | 0.29 | 0.29 | 0.26 | 0.3491 | 0.33 | 0.4092 | 0.29709 | 2.63e-05 | 0.41231 | 0.35 | 0.383 | 0.31 | 0.3159 | 0.5 | 0.35 | 0.26 | 0.29 | 0.31 | 0.31 | 0.37 | 0.33 |  |  | 1
| 1 | 2 | 0.29 | 0.27 | 0.3 | 0.28 | 0.32 | 0.29 | 0.29 | 0.3 | 0.26 | 0.3363 | 0.34 | 0.382 | 0.31183 | 2.78e-05 | 0.40318 | 0.34 | 0.38 | 0.3 | 0.3124 | 0.5 | 0.33 | 0.25 | 0.29 | 0.3 | 0.31 | 0.36 | 0.27 |  |  | 0
| 1 | 3 | 0.32 | 0.3 | 0.3 | 0.28 | 0.36 | 0.3 | 0.28 | 0.3 | 0.27 | 0.3262 | 0.33 | 0.3961 | 0.30221 | 2.54e-05 | 0.40842 | 0.35 | 0.385 | 0.31 | 0.3094 | 0.49 | 0.34 | 0.24 | 0.28 | 0.31 | 0.32 | 0.36 | 0.28 |  |  | 16
| 1 | 4 | 0.32 | 0.29 | 0.31 | 0.27 | 0.34 | 0.3 | 0.29 | 0.3 | 0.26 | 0.3186 | 0.32 | 0.3905 | 0.31024 | 2.62e-05 | 0.39433 | 0.32 | 0.385 | 0.31 | 0.3138 | 0.51 | 0.34 | 0.25 | 0.28 | 0.31 | 0.33 | 0.36 | 0.32 |  |  | 0.31773
| 1 | 5 | 0.33 | 0.27 | 0.31 | 0.27 | 0.34 | 0.3 | 0.29 | 0.3 | 0.26 | 0.3205 | 0.34 | 0.3972 | 0.30561 | 2.53e-05 | 0.38862 | 0.37 | 0.378 | 0.32 | 0.3142 | 0.51 | 0.33 | 0.25 | 0.29 | 0.31 | 0.32 | 0.36 | 0.27 |  |  | 88
| 1 | 6 | 0.32 | 0.28 | 0.3 | 0.28 | 0.34 | 0.31 | 0.28 | 0.29 | 0.26 | 0.2845 | 0.33 | 0.4193 | 0.29841 | 2.7e-05 | 0.40084 | 0.38 | 0.386 | 0.31 | 0.3127 | 0.49 | 0.33 | 0.25 | 0.29 | 0.31 | 0.32 | 0.36 | 0.32
| Mean |  | 0.31667 | 0.28333 | 0.3 | 0.275 | 0.34333 | 0.29833 | 0.28667 | 0.29667 | 0.26167 | 0.32253 | 0.33167 | 0.39905 | 0.30423 | 2.6333e-05 | 0.40128 | 0.35167 | 0.38283 | 0.31 | 0.31307 | 0.5 | 0.33667 | 0.25 | 0.28667 | 0.30833 | 0.31833 | 0.36167 | 0.29833
| Median |  | 0.32 | 0.285 | 0.3 | 0.275 | 0.34 | 0.3 | 0.29 | 0.3 | 0.26 | 0.32335 | 0.33 | 0.39665 | 0.30391 | 2.625e-05 | 0.40201 | 0.35 | 0.384 | 0.31 | 0.31325 | 0.5 | 0.335 | 0.25 | 0.29 | 0.31 | 0.32 | 0.36 | 0.3
| Std Dev. |  | 0.013663 | 0.012111 | 0.010954 | 0.0054772 | 0.015055 | 0.0075277 | 0.005164 | 0.005164 | 0.0040825 | 0.021794 | 0.0075277 | 0.013334 | 0.0060779 | 9.5429e-07 | 0.0087764 | 0.02137 | 0.0031885 | 0.0063246 | 0.0021851 | 0.0089443 | 0.008165 | 0.0063246 | 0.005164 | 0.0040825 | 0.0075277 | 0.0040825 | 0.027869
| Rel.Std.Dev. |  | 4.315% | 4.274% | 3.651% | 1.992% | 4.385% | 2.523% | 1.801% | 1.741% | 1.56% | 6.757% | 2.27% | 3.342% | 1.998% | 3.624% | 2.187% | 6.077% | 0.8329% | 2.04% | 0.698% | 1.789% | 2.425% | 2.53% | 1.801% | 1.324% | 2.365% | 1.129% | 9.341%
| PDM3 |  | -0.3357% | -10.83% | -5.581% | -13.45% | 8.057% | -6.106% | -9.778% | -6.63% | -17.65% | 1.511% | 4.385% | 25.59% | -4.249% | -99.99% | 26.3% | 10.68% | 20.49% | -2.434% | -1.469% | 57.36% | 5.959% | -21.32% | -9.778% | -2.958% | 0.1888% | 13.83% | -6.106%
| Z-Score (Absolute) |  | 0.14 | 0.76 | 0.31 | 0.99 | 0.86 | 0.36 | 0.67 | 0.4 | 1.35 | 0.3 | 0.55 | 2.37 | 0.2 | NA | 2.43 | 1.09 | 1.93 | 0.04 | 0.04 | 5.11 | 0.68 | 1.67 | 0.67 | 0.09 | 0.18 | 1.36 | 0.36
| Analytical results for La in OREAS 627 (Certified Value 15.6 ppm) | [30 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 1
| 1 | 1 | 11 | 9.6 | 13 | 13.9 | 16.3 | 13.6 | 15.6 | 14.6 | 13.8 | 15.571 | 13.8 | 23 | 15.6893 | 19.1674 | 16.6 | 24.9 | 10 | 17.28 | 24.5 | 16.8 | 16 | 14.4 | 14.4 | 16.1 | 19.4 | 18.4 |  |  |  | 1
| 1 | 2 | 11 | 9.1 | 14 | 13.9 | 14.9 | 14.2 | 14.9 | 15.5 | 13.9 | 15.244 | 13.6 | 21 | 16.3875 | 19.149 | 16 | 23.8 | 10 | 17.84 | 24.4 | 17.2 | 16 | 14.7 | 13.9 | 15.8 | 19.3 | 16.9 |  |  |  | 34
| 1 | 3 | 11 | 9 | 14 | 13.9 | 16.8 | 14.2 | 14.2 | 14.7 | 14.3 | 14.812 | 13 | 22 | 16.2497 | 19.0179 | 15.8 | 24.3 | 10 | 18.02 | 23.8 | 16.8 | 16 | 15.5 | 14.4 | 15.5 | 19.2 | 17.6 |  |  |  | 16
| 1 | 4 | 11 | 9.8 | 14 | 13.5 | 15.4 | 13.8 | 14.9 | 14.4 | 13.7 | 14.455 | 13.2 | 22 | 16.6104 | 18.8998 | 14.8 | 24 | 10 | 17.65 | 24.1 | 16.4 | 16 | 15.3 | 14.8 | 14.8 | 19 | 17.6 |  |  |  | 15.6447
| 1 | 5 | 11 | 8.5 | 15 | 13.6 | 15.8 | 13.8 | 14.7 | 14.3 | 14.3 | 14.793 | 14.3 | 22 | 16.5579 | 18.7986 | 16.4 | 24.3 | 10 | 17.84 | 24.1 | 16.8 | 16 | 14 | 13.8 | 14.5 | 19 | 16 |  |  |  | 89
| 1 | 6 | 11 | 8.3 | 14 | 13.9 | 14.7 | 13.7 | 14.8 | 15.1 | 14.4 | 13.937 | 12.6 | 23 | 16.2993 | 18.7252 | 16.7 | 24 | 10 | 17.63 | 23.5 | 15.4 | 17 | 14.2 | 14.4 | 14.7 | 18.5 | 17.3
| Mean |  | 11 | 9.05 | 14 | 13.7833 | 15.65 | 13.8833 | 14.85 | 14.7667 | 14.0667 | 14.802 | 13.4167 | 22.1667 | 16.299 | 18.9596 | 16.05 | 24.2167 | 10 | 17.71 | 24.0667 | 16.5667 | 16.1667 | 14.6833 | 14.2833 | 15.2333 | 19.0667 | 17.3
| Median |  | 11 | 9.05 | 14 | 13.9 | 15.6 | 13.8 | 14.85 | 14.65 | 14.1 | 14.8025 | 13.4 | 22 | 16.3434 | 18.9589 | 16.2 | 24.15 | 10 | 17.745 | 24.1 | 16.8 | 16 | 14.55 | 14.4 | 15.15 | 19.1 | 17.45
| Std Dev. |  | 0 | 0.58907 | 0.63246 | 0.18348 | 0.81179 | 0.25626 | 0.45056 | 0.45461 | 0.30111 | 0.57542 | 0.608 | 0.75277 | 0.33045 | 0.18272 | 0.70356 | 0.38687 | 0 | 0.25487 | 0.37238 | 0.62503 | 0.40825 | 0.6047 | 0.37103 | 0.65625 | 0.32042 | 0.80498
| Rel.Std.Dev. |  | 0% | 6.509% | 4.518% | 1.331% | 5.187% | 1.846% | 3.034% | 3.079% | 2.141% | 3.887% | 4.532% | 3.396% | 2.027% | 0.9637% | 4.384% | 1.598% | 0% | 1.439% | 1.547% | 3.773% | 2.525% | 4.118% | 2.598% | 4.308% | 1.681% | 4.653%
| PDM3 |  | -29.69% | -42.15% | -10.51% | -11.9% | 0.03408% | -11.26% | -5.079% | -5.612% | -10.09% | -5.386% | -14.24% | 41.69% | 4.183% | 21.19% | 2.591% | 54.79% | -36.08% | 13.2% | 53.83% | 5.893% | 3.337% | -6.145% | -8.702% | -2.629% | 21.87% | 10.58%
| Z-Score (Absolute) |  | NA | 3.13 | NA | 0.73 | 0.21 | 0.68 | 0.19 | 0.24 | 0.59 | 0.22 | 0.92 | NA | 0.54 | 1.88 | 0.41 | 4.54 | NA | 1.25 | 4.47 | 0.67 | NA | 0.28 | 0.48 | 0 | 1.94 | 1.05
Results from laboratories 1.01, 1.03, 1.12, 1.18 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Li in OREAS 627 (Certified Value 5.18 ppm) | [30 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.13 | 1.14 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 0.2g | 0.5g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [6 empty] | 2
| 1 | 1 | 5 | 6.5 | 4.5 | 6.2 | 5.3 | 5.3 | 5.1 | 3.7 | 5.7 | 7 | 4.7509 | 4.99 | 5.4329 | 6.5 | 5.3 | 7.04 | 7 | 7 | 5 | 5 | 5.2 | 9 | 5 | [6 empty] | 1
| 1 | 2 | 5 | 5.3 | 4.6 | 5.7 | 5.3 | 4.9 | 5.7 | 4 | 5.6 | 6 | 5.277 | 2.55 | 5.2938 | 6 | 5 | 7.08 | 6.7 | 7 | 6 | 5 | 4.7 | 5 | 5 | [6 empty] | 0
| 1 | 3 | 5 | 5.8 | 4.7 | 6.3 | 5.4 | 4.7 | 5.7 | 3.9 | 5.8 | 6 | 5.2348 | 4.92 | 5.1733 | 6.2 | 4 | 7.09 | 7 | 7 | 6 | 5 | 4.8 | 3 | 5 | [6 empty] | 16
| 1 | 4 | 6 | 6 | 4.8 | 5.4 | 5.6 | 4.8 | 5.8 | 3.8 | 5.3 | 6 | 5.216 | 1.96 | 4.6868 | 6.3 | 5.6 | 7.11 | 7.3 | 7 | 5 | 6 | 4.6 | 6 | 5 | [6 empty] | 5.1819
| 1 | 5 | 6 | 5.3 | 4.7 | 5.5 | 5.8 | 4.9 | 5.7 | 4.2 | 5.2 | 6 | 5.3166 | 4.75 | 4.8443 | 6.2 | 5.5 | 7.12 | 7.2 | 7 | 5 | 6 | 4.8 | 8 | 5 | [6 empty] | 90
| 1 | 6 | 6 | 5.4 | 5.3 | 5.5 | 5.8 | 4.8 | 5.1 | 4.5 | 4.8 | 7 | 4.8766 | 2.43 | 4.5798 | 6.2 | 4.9 | 7.15 | 7.1 | 7 | 5 | 6 | 4.9 | < 1 | 5
| Mean |  | 5.5 | 5.7167 | 4.7667 | 5.7667 | 5.5333 | 4.9 | 5.5167 | 4.0167 | 5.4 | 6.3333 | 5.112 | 3.6 | 5.0018 | 6.2333 | 5.05 | 7.0983 | 7.05 | 7 | 5.3333 | 5.5 | 4.8333 | 6.2 | 5
| Median |  | 5.5 | 5.6 | 4.7 | 5.6 | 5.5 | 4.85 | 5.7 | 3.95 | 5.45 | 6 | 5.2254 | 3.65 | 5.0088 | 6.2 | 5.15 | 7.1 | 7.05 | 7 | 5 | 5.5 | 4.8 | 6 | 5
| Std Dev. |  | 0.54772 | 0.47924 | 0.28048 | 0.38816 | 0.23381 | 0.20976 | 0.32506 | 0.29269 | 0.37417 | 0.5164 | 0.23696 | 1.4253 | 0.34718 | 0.1633 | 0.58224 | 0.037639 | 0.20736 | 0 | 0.5164 | 0.54772 | 0.20656 | 2.3875 | 0
| Rel.Std.Dev. |  | 9.959% | 8.383% | 5.884% | 6.731% | 4.225% | 4.281% | 5.892% | 7.287% | 6.929% | 8.154% | 4.635% | 39.59% | 6.941% | 2.62% | 11.53% | 0.5302% | 2.941% | 0% | 9.682% | 9.959% | 4.274% | 38.51% | 0%
| PDM3 |  | 6.138% | 10.32% | -8.014% | 11.28% | 6.781% | -5.441% | 6.46% | -22.49% | 4.208% | 22.22% | -1.35% | -30.53% | -3.476% | 20.29% | -2.546% | 36.98% | 36.05% | 35.08% | 2.922% | 6.138% | -6.727% | 19.65% | -3.511%
| Z-Score (Absolute) |  | NA | 0.7 | 0.75 | 0.78 | 0.42 | 0.54 | 0.4 | 1.89 | 0.22 | NA | 0.22 | 2.53 | 0.39 | 1.49 | 0.31 | 2.81 | 2.74 | NA | NA | NA | 0.65 | NA | NA
Results from laboratories 1.01, 1.12, 1.23, 1.25, 1.26, 1.28 and 1.30 were removed due to their 1 ppm reading resolution.
| Analytical results for Lu in OREAS 627 (Certified Value 0.064 ppm) | [30 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.23; 1.05: 1.25; 1.06: 1.26 |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*OES/MS; 1.06: AR*OES/MS |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.25g; 1.05: 0.25g; 1.06: 0.25g |  |  |  | 3
| 1 | 1 | 1.01: < 0.1; 1.02: 0.075; 1.03: 0.061596; 1.04: 0.07; 1.05: 0.07; 1.06: 0.05 |  |  |  | 1
| 1 | 2 | 1.01: < 0.1; 1.02: 0.071; 1.03: 0.058823; 1.04: 0.08; 1.05: 0.07; 1.06: 0.05 |  |  |  | 35
| 1 | 3 | 1.01: < 0.1; 1.02: 0.066; 1.03: 0.057338; 1.04: 0.07; 1.05: 0.07; 1.06: 0.05 |  |  |  | 16
| 1 | 4 | 1.01: < 0.1; 1.02: 0.072; 1.03: 0.060865; 1.04: 0.07; 1.05: 0.07; 1.06: 0.05 |  |  |  | 0.063728
| 1 | 5 | 1.01: < 0.1; 1.02: 0.067; 1.03: 0.055409; 1.04: 0.08; 1.05: 0.07; 1.06: 0.05 |  |  |  | 91
| 1 | 6 | 1.01: < 0.1; 1.02: 0.058; 1.03: < 0.05; 1.04: 0.07; 1.05: 0.06; 1.06: 0.05
| Mean |  | 1.02: 0.068167; 1.03: 0.058806; 1.04: 0.073333; 1.05: 0.068333; 1.06: 0.05
| Median |  | 1.02: 0.069; 1.03: 0.058823; 1.04: 0.07; 1.05: 0.07; 1.06: 0.05
| Std Dev. |  | 1.02: 0.0059805; 1.03: 0.0025357; 1.04: 0.005164; 1.05: 0.0040825; 1.06: 7.6012e-18
| Rel.Std.Dev. |  | 1.02: 8.773%; 1.03: 4.312%; 1.04: 7.042%; 1.05: 5.974%; 1.06: 1.52e-14%
| PDM3 |  | 1.02: 6.965%; 1.03: -7.723%; 1.04: 15.07%; 1.05: 7.227%; 1.06: -21.54%
| Z-Score (Absolute) |  | 1.01: 0.78; 1.02: 0.55; 1.03: 0.55; 1.04: 0.93; 1.05: 0.56; 1.06: 0.78
| Analytical results for Mg in OREAS 627 (Certified Value 0.231 wt.%) | [30 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.23 | 0.24 | 0.22 | 0.21 | 0.23 | 0.2 | 0.22 | 0.22 | 0.24 | 0.25 | 0.22 | 0.25 | 0.22014 | 0.328 | 0.25613 | 0.22 | 0.25 | 0.23 | 0.2392 | 0.25 | 0.24 | 0.22 | 0.22 | 0.23 | 0.22 | 0.238 | 0.26 |  |  | 1
| 1 | 2 | 0.21 | 0.2 | 0.23 | 0.21 | 0.22 | 0.21 | 0.21 | 0.23 | 0.24 | 0.24 | 0.23 | 0.23 | 0.23444 | 0.316 | 0.24731 | 0.22 | 0.24 | 0.23 | 0.2365 | 0.25 | 0.24 | 0.21 | 0.23 | 0.23 | 0.21 | 0.242 | 0.23 |  |  | 0
| 1 | 3 | 0.24 | 0.21 | 0.24 | 0.21 | 0.24 | 0.21 | 0.21 | 0.22 | 0.25 | 0.24 | 0.22 | 0.24 | 0.23304 | 0.353 | 0.25353 | 0.22 | 0.25 | 0.24 | 0.2346 | 0.25 | 0.24 | 0.2 | 0.22 | 0.24 | 0.22 | 0.244 | 0.24 |  |  | 16
| 1 | 4 | 0.24 | 0.22 | 0.24 | 0.21 | 0.22 | 0.22 | 0.21 | 0.23 | 0.25 | 0.23 | 0.23 | 0.24 | 0.23493 | 0.391 | 0.24272 | 0.22 | 0.24 | 0.24 | 0.2374 | 0.26 | 0.24 | 0.21 | 0.22 | 0.24 | 0.23 | 0.243 | 0.24 |  |  | 0.23061
| 1 | 5 | 0.25 | 0.2 | 0.24 | 0.21 | 0.22 | 0.21 | 0.21 | 0.22 | 0.25 | 0.23 | 0.23 | 0.23 | 0.23481 | 0.375 | 0.24228 | 0.23 | 0.24 | 0.25 | 0.2377 | 0.26 | 0.24 | 0.21 | 0.22 | 0.24 | 0.21 | 0.246 | 0.22 |  |  | 92
| 1 | 6 | 0.24 | 0.21 | 0.23 | 0.22 | 0.22 | 0.22 | 0.21 | 0.21 | 0.25 | 0.2 | 0.23 | 0.24 | 0.22802 | 0.33 | 0.24352 | 0.24 | 0.24 | 0.23 | 0.2383 | 0.25 | 0.23 | 0.22 | 0.22 | 0.24 | 0.22 | 0.234 | 0.25
| Mean |  | 0.235 | 0.21333 | 0.23333 | 0.21167 | 0.225 | 0.21167 | 0.21167 | 0.22167 | 0.24667 | 0.23167 | 0.22667 | 0.23833 | 0.2309 | 0.34883 | 0.24758 | 0.225 | 0.24333 | 0.23667 | 0.23728 | 0.25333 | 0.23833 | 0.21167 | 0.22167 | 0.23667 | 0.21833 | 0.24117 | 0.24
| Median |  | 0.24 | 0.21 | 0.235 | 0.21 | 0.22 | 0.21 | 0.21 | 0.22 | 0.25 | 0.235 | 0.23 | 0.24 | 0.23374 | 0.3415 | 0.24541 | 0.22 | 0.24 | 0.235 | 0.23755 | 0.25 | 0.24 | 0.21 | 0.22 | 0.24 | 0.22 | 0.2425 | 0.24
| Std Dev. |  | 0.013784 | 0.015055 | 0.008165 | 0.0040825 | 0.0083666 | 0.0075277 | 0.0040825 | 0.0075277 | 0.005164 | 0.017224 | 0.005164 | 0.0075277 | 0.005877 | 0.029485 | 0.0059474 | 0.0083666 | 0.005164 | 0.008165 | 0.0015943 | 0.005164 | 0.0040825 | 0.0075277 | 0.0040825 | 0.005164 | 0.0075277 | 0.0044008 | 0.014142
| Rel.Std.Dev. |  | 5.866% | 7.057% | 3.499% | 1.929% | 3.718% | 3.556% | 1.929% | 3.396% | 2.094% | 7.435% | 2.278% | 3.158% | 2.545% | 8.452% | 2.402% | 3.718% | 2.122% | 3.45% | 0.6719% | 2.038% | 1.713% | 3.556% | 1.842% | 2.182% | 3.448% | 1.825% | 5.893%
| PDM3 |  | 1.906% | -7.49% | 1.183% | -8.213% | -2.431% | -8.213% | -8.213% | -3.876% | 6.965% | 0.4602% | -1.708% | 3.351% | 0.1259% | 51.27% | 7.361% | -2.431% | 5.519% | 2.628% | 2.896% | 9.856% | 3.351% | -8.213% | -3.876% | 2.628% | -5.322% | 4.58% | 4.074%
| Z-Score (Absolute) |  | 0.13 | 1.62 | 0 | 1.75 | 0.67 | 1.75 | 1.75 | 0.94 | 1.08 | 0.13 | 0.54 | 0.4 | 0.2 | 9.35 | 1.15 | 0.67 | 0.81 | 0.27 | 0.32 | 1.62 | 0.4 | 1.75 | 0.94 | 0.27 | 1.21 | 0.63 | 0.54
| Analytical results for Mn in OREAS 627 (Certified Value 0.644 wt.%) | [30 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.3776 | 0.588 | 0.4579 | 0.63 | 0.679 | 0.65 | 0.645 | 0.649 | 0.568 | 0.6907 | 0.635 | 0.679 | 0.073541 | 0.534 | 0.63273 | 0.6366 | 0.67 | 0.5749 | 0.6587 | 0.6797 | 0.6829 | 0.6819 | 0.6229 | 0.6143 | 0.6242 | 0.6861 | 0.6753 |  |  | 1
| 1 | 2 | 0.3711 | 0.565 | 0.4595 | 0.626 | 0.631 | 0.662 | 0.63 | 0.657 | 0.573 | 0.6711 | 0.6505 | 0.6547 | 0.078324 | 0.635 | 0.62987 | 0.6301 | 0.675 | 0.582 | 0.6652 | 0.6829 | 0.665 | 0.6786 | 0.6258 | 0.6591 | 0.6181 | 0.6685 | 0.6698 |  |  | 21
| 1 | 3 | 0.3756 | 0.59 | 0.4772 | 0.635 | 0.696 | 0.667 | 0.622 | 0.645 | 0.6 | 0.6477 | 0.6181 | 0.6737 | 0.077857 | 0.634 | 0.63279 | 0.6354 | 0.677 | 0.5877 | 0.6741 | 0.6655 | 0.6774 | 0.6656 | 0.628 | 0.6442 | 0.6383 | 0.6746 | 0.6653 |  |  | 16
| 1 | 4 | 0.3786 | 0.583 | 0.4681 | 0.634 | 0.661 | 0.673 | 0.628 | 0.653 | 0.576 | 0.6461 | 0.6244 | 0.6777 | 0.078487 | 0.723 | 0.62904 | 0.621 | 0.682 | 0.5765 | 0.6689 | 0.6586 | 0.6699 | 0.6851 | 0.6251 | 0.6465 | 0.6558 | 0.6797 | 0.6715 |  |  | 0.64421
| 1 | 5 | 0.3793 | 0.56 | 0.4741 | 0.642 | 0.664 | 0.655 | 0.642 | 0.646 | 0.585 | 0.6453 | 0.654 | 0.6642 | 0.078448 | 0.62 | 0.63176 | 0.6319 | 0.671 | 0.598 | 0.6721 | 0.663 | 0.6681 | 0.6715 | 0.6243 | 0.6551 | 0.6452 | 0.6712 | 0.6609 |  |  | 93
| 1 | 6 | 0.3798 | 0.561 | 0.4764 | 0.636 | 0.657 | 0.671 | 0.623 | 0.643 | 0.566 | 0.5794 | 0.6071 | 0.6848 | 0.07617 | 0.614 | 0.62765 | 0.6349 | 0.674 | 0.5852 | 0.6737 | 0.6702 | 0.6765 | 0.6819 | 0.6272 | 0.6494 | 0.6552 | 0.6768 | 0.665
| Mean |  | 0.377 | 0.5745 | 0.46887 | 0.63383 | 0.66467 | 0.663 | 0.63167 | 0.64883 | 0.578 | 0.64672 | 0.63152 | 0.67235 | 0.077138 | 0.62667 | 0.63064 | 0.63165 | 0.67483 | 0.58405 | 0.66878 | 0.66998 | 0.6733 | 0.67743 | 0.62555 | 0.64477 | 0.63947 | 0.67615 | 0.66797
| Median |  | 0.3781 | 0.574 | 0.4711 | 0.6345 | 0.6625 | 0.6645 | 0.629 | 0.6475 | 0.5745 | 0.6469 | 0.6297 | 0.6757 | 0.07809 | 0.627 | 0.63081 | 0.6334 | 0.6745 | 0.5836 | 0.6705 | 0.66785 | 0.6732 | 0.68025 | 0.62545 | 0.64795 | 0.64175 | 0.6757 | 0.66755
| Std Dev. |  | 0.0032502 | 0.013982 | 0.0085109 | 0.0054559 | 0.021878 | 0.0090995 | 0.0096885 | 0.0053072 | 0.012696 | 0.037587 | 0.018465 | 0.011024 | 0.0019665 | 0.060325 | 0.002112 | 0.0057448 | 0.0043551 | 0.0084164 | 0.005966 | 0.009589 | 0.0067335 | 0.0074199 | 0.0018748 | 0.015907 | 0.015736 | 0.0062842 | 0.0052053
| Rel.Std.Dev. |  | 0.8621% | 2.434% | 1.815% | 0.8608% | 3.292% | 1.372% | 1.534% | 0.818% | 2.197% | 5.812% | 2.924% | 1.64% | 2.549% | 9.626% | 0.3349% | 0.9095% | 0.6454% | 1.441% | 0.8921% | 1.431% | 1% | 1.095% | 0.2997% | 2.467% | 2.461% | 0.9294% | 0.7793%
| PDM3 |  | -41.48% | -10.82% | -27.22% | -1.61% | 3.176% | 2.917% | -1.946% | 0.7183% | -10.28% | 0.3897% | -1.97% | 4.369% | -88.03% | -2.723% | -2.106% | -1.949% | 4.754% | -9.338% | 3.815% | 4.001% | 4.516% | 5.158% | -2.896% | 0.08701% | -0.7357% | 4.959% | 3.688%
| Z-Score (Absolute) |  | 6.21 | 1.54 | 4.04 | 0.13 | 0.6 | 0.56 | 0.18 | 0.22 | 1.45 | 0.17 | 0.19 | 0.78 | 13.3 | 0.3 | 0.21 | 0.18 | 0.84 | 1.31 | 0.69 | 0.72 | 0.8 | 0.9 | 0.33 | 0.13 | 0 | 0.87 | 0.67
| Analytical results for Mo in OREAS 627 (Certified Value 28.7 ppm) | [30 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 2
| 1 | 1 | 26 | 21.2 | 29.3 | 28.1 | 29.8 | 26 | 28.2 | 29.4 | 29.1 | 31.4 | 28.53 | 30 | 27.51 | 30.586 | 27.552 | 28.1 | 29.5 | 26.8 | 27.2 | 28.2 | 32.36 | 30.67 | 31 | 28.8 | 31.25 | 25.46 | 30 | 28.19 |  | 1
| 1 | 2 | 25 | 19.1 | 28.9 | 28.3 | 28 | 28 | 30.3 | 29.4 | 29.3 | 31 | 28.48 | 30 | 29.0859 | 29.478 | 26.8923 | 28.3 | 30.2 | 26.8 | 27.15 | 28.6 | 31.59 | 30.13 | 30 | 31 | 31.31 | 25.76 | 32 | 23.99 |  | 37
| 1 | 3 | 27 | 13.5 | 29.7 | 27.3 | 31.7 | 25.8 | 26.4 | 29.4 | 29.9 | 30.5 | 27.5 | 30 | 28.2468 | 29.269 | 27.0012 | 26.9 | 30 | 27 | 27.06 | 28 | 31.3 | 29.73 | 30 | 31.1 | 31.65 | 25.7 | 35 | 26.63 |  | 16
| 1 | 4 | 27 | 20.3 | 29.4 | 28.9 | 27.8 | 24 | 25.5 | 28.3 | 29 | 30.8 | 27.97 | 29 | 28.2637 | 29.857 | 27.2225 | 25.6 | 30.7 | 26.8 | 27.19 | 28.6 | 29.44 | 29.32 | 31 | 28.1 | 31.23 | 24.14 | 32 | 27.12 |  | 28.7071
| 1 | 5 | 27 | 19 | 29.6 | 28.8 | 29.9 | 25 | 25.4 | 28.9 | 30.1 | 30.4 | 28.86 | 30 | 28.2684 | 30.368 | 27.7574 | 26.3 | 30.6 | 27.3 | 27.14 | 28.2 | 31.91 | 29.83 | 30 | 30.8 | 31.71 | 26.07 | 33 | 34.07 |  | 94
| 1 | 6 | 27 | 21.9 | 29.4 | 29 | 29.1 | 25.2 | 26.3 | 30.1 | 29.6 | 29.4 | 27.98 | 30 | 28.7189 | 27.727 | 26.7135 | 28.7 | 30.2 | 27 | 27.19 | 28.3 | 31.45 | 30.6 | 31 | 28.4 | 31.36 | 26.13 | 30 | 25.95
| Mean |  | 26.5 | 19.1667 | 29.3833 | 28.4 | 29.3833 | 25.6667 | 27.0167 | 29.25 | 29.5 | 30.5833 | 28.22 | 29.8333 | 28.349 | 29.5475 | 27.1898 | 27.3167 | 30.2 | 26.95 | 27.155 | 28.3167 | 31.3417 | 30.0467 | 30.5 | 29.7 | 31.4183 | 25.5433 | 32 | 27.6583
| Median |  | 27 | 19.7 | 29.4 | 28.55 | 29.45 | 25.5 | 26.35 | 29.4 | 29.45 | 30.65 | 28.23 | 30 | 28.266 | 29.6675 | 27.1118 | 27.5 | 30.2 | 26.9 | 27.17 | 28.25 | 31.52 | 29.98 | 30.5 | 29.8 | 31.335 | 25.73 | 32 | 26.875
| Std Dev. |  | 0.83666 | 3.0011 | 0.27869 | 0.64498 | 1.4359 | 1.3426 | 1.8968 | 0.60249 | 0.44272 | 0.6824 | 0.4923 | 0.40825 | 0.53064 | 1.0241 | 0.40134 | 1.2368 | 0.43359 | 0.19748 | 0.05244 | 0.24014 | 1.0051 | 0.52477 | 0.54772 | 1.4085 | 0.20865 | 0.73066 | 1.8974 | 3.4384
| Rel.Std.Dev. |  | 3.157% | 15.66% | 0.9485% | 2.271% | 4.887% | 5.231% | 7.021% | 2.06% | 1.501% | 2.231% | 1.745% | 1.368% | 1.872% | 3.466% | 1.476% | 4.528% | 1.436% | 0.7328% | 0.1931% | 0.848% | 3.207% | 1.747% | 1.796% | 4.743% | 0.6641% | 2.86% | 5.929% | 12.43%
| PDM3 |  | -7.688% | -33.23% | 2.356% | -1.07% | 2.356% | -10.59% | -5.889% | 1.891% | 2.762% | 6.536% | -1.697% | 3.923% | -1.248% | 2.928% | -5.285% | -4.844% | 5.2% | -6.121% | -5.407% | -1.36% | 9.177% | 4.666% | 6.246% | 3.459% | 9.444% | -11.02% | 11.47% | -3.653%
| Z-Score (Absolute) |  | 1.09 | 4.52 | 0.26 | 0.2 | 0.26 | 1.48 | 0.85 | 0.2 | 0.32 | 0.82 | 0.28 | 0.47 | 0.22 | 0.34 | 0.76 | 0.71 | 0.64 | 0.88 | 0.78 | 0.24 | 1.18 | 0.57 | 0.78 | 0.41 | 1.21 | 1.53 | 1.49 | 0.55
| Analytical results for Na in OREAS 627 (Certified Value 0.041 wt.%) | [30 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 0.04 | 0.038 | 0.034 | 0.03 | 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.047 | 0.04 | 0.04258 | < 0.1 | 0.037546 | 0.06 | 0.04 | 0.03 | 0.0465 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.052 | 0.05 |  |  | 1
| 1 | 2 | 0.04 | 0.034 | 0.036 | 0.03 | 0.04 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.049 | 0.04 | 0.043156 | < 0.1 | 0.035278 | 0.06 | 0.04 | 0.03 | 0.0466 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.053 | 0.05 |  |  | 13
| 1 | 3 | 0.04 | 0.036 | 0.036 | 0.03 | 0.06 | 0.04 | 0.03 | 0.04 | 0.03 | 0.04 | 0.047 | 0.04 | 0.043062 | < 0.1 | 0.033675 | 0.06 | 0.04 | 0.03 | 0.0468 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.051 | 0.05 |  |  | 16
| 1 | 4 | 0.04 | 0.036 | 0.037 | 0.03 | 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.047 | 0.04 | 0.043796 | < 0.1 | 0.032285 | 0.06 | 0.04 | 0.04 | 0.0468 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 |  |  | 0.041092
| 1 | 5 | 0.04 | 0.034 | 0.037 | 0.03 | 0.05 | 0.04 | 0.03 | 0.04 | 0.04 | 0.03 | 0.05 | 0.04 | 0.043823 | < 0.1 | 0.033694 | 0.07 | 0.04 | 0.04 | 0.0472 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.05 | 0.05 |  |  | 95
| 1 | 6 | 0.04 | 0.034 | 0.036 | 0.03 | 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.03 | 0.05 | 0.04 | 0.042696 | < 0.1 | 0.036205 | 0.06 | 0.04 | 0.04 | 0.0471 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.051 | 0.05
| Mean |  | 0.04 | 0.035333 | 0.036 | 0.03 | 0.05 | 0.04 | 0.03 | 0.04 | 0.031667 | 0.035 | 0.048333 | 0.04 | 0.043186 |  | 0.034781 | 0.061667 | 0.04 | 0.035 | 0.046833 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.051167 | 0.05
| Median |  | 0.04 | 0.035 | 0.036 | 0.03 | 0.05 | 0.04 | 0.03 | 0.04 | 0.03 | 0.035 | 0.048 | 0.04 | 0.043109 |  | 0.034486 | 0.06 | 0.04 | 0.035 | 0.0468 | 0.05 | 0.05 | 0.05 | 0.04 | 0.04 | 0.04 | 0.051 | 0.05
| Std Dev. |  | 0 | 0.001633 | 0.0010954 | 0 | 0.0063246 | 0 | 0 | 0 | 0.0040825 | 0.0054772 | 0.0015055 | 0 | 0.00052927 |  | 0.0019262 | 0.0040825 | 0 | 0.0054772 | 0.00027325 | 7.6012e-18 | 7.6012e-18 | 7.6012e-18 | 0 | 0 | 0 | 0.001169 | 7.6012e-18
| Rel.Std.Dev. |  | 0% | 4.622% | 3.043% | 0% | 12.65% | 0% | 0% | 0% | 12.89% | 15.65% | 3.115% | 0% | 1.226% |  | 5.538% | 6.62% | 0% | 15.65% | 0.5835% | 1.52e-14% | 1.52e-14% | 1.52e-14% | 0% | 0% | 0% | 2.285% | 1.52e-14%
| PDM3 |  | -2.657% | -14.01% | -12.39% | -26.99% | 21.68% | -2.657% | -26.99% | -2.657% | -22.94% | -14.83% | 17.62% | -2.657% | 5.095% |  | -15.36% | 50.07% | -2.657% | -14.83% | 13.97% | 21.68% | 21.68% | 21.68% | -2.657% | -2.657% | -2.657% | 24.52% | 21.68%
| Z-Score (Absolute) |  | 0 | 0.6 | 0.52 | 1.29 | 1.29 | 0 | 1.29 | 0 | 1.08 | 0.65 | 1.08 | 0 | 0.41 | 1.29 | 0.67 | 2.8 | 0 | 0.65 | 0.88 | 1.29 | 1.29 | 1.29 | 0 | 0 | 0 | 1.44 | 1.29
| Analytical results for Nb in OREAS 627 (Certified Value 0.79 ppm) | [30 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 2
| 1 | 1 | < 5 | 1.1 | 0.59 | 0.93 | 0.76 | 0.66 | 0.6 | 0.89 | 1.36 | 0.7902 | 1.1102 | 1.1 | 0.3 | 0.86 | 0.61 | 0.79 | 0.83 | 1.37 | [11 empty] | 1
| 1 | 2 | < 5 | 1.2 | 0.58 | 0.82 | 0.77 | 0.68 | 0.6 | 0.84 | 1.33 | 0.74267 | 1.0427 | 1 | 0.4 | 0.83 | 0.63 | 0.85 | 0.92 | 0.68 | [11 empty] | 38
| 1 | 3 | < 5 | 1.3 | 0.59 | 0.93 | 0.77 | 0.64 | 0.62 | 0.89 | 1.29 | 0.6969 | 1.0366 | 1.1 | 0.3 | 0.76 | 0.65 | 0.81 | 0.84 | 0.82 | [11 empty] | 16
| 1 | 4 | < 5 | 1.2 | 0.57 | 0.85 | 0.71 | 0.66 | 0.62 | 0.83 | 1.31 | 0.76534 | 1.0251 | 1.1 | 0.3 | 0.8 | 0.69 | 0.75 | 0.78 | 1.28 | [11 empty] | 0.7903
| 1 | 5 | < 5 | 1.1 | 0.59 | 0.92 | 0.73 | 0.63 | 0.59 | 0.92 | 1.34 | 0.72346 | 1.114 | 1.1 | 0.3 | 0.81 | 0.68 | 0.78 | 0.71 | 0.9 | [11 empty] | 96
| 1 | 6 | < 5 | 1.2 | 0.57 | 0.88 | 0.74 | 0.64 | 0.64 | 0.86 | 1.29 | 0.75931 | 1.0571 | 1.1 | 0.3 | 0.79 | 0.56 | 0.8 | 1.08 | 1.81
| Mean |  |  | 1.1833 | 0.58167 | 0.88833 | 0.74667 | 0.65167 | 0.61167 | 0.87167 | 1.32 | 0.74631 | 1.0643 | 1.0833 | 0.31667 | 0.80833 | 0.63667 | 0.79667 | 0.86 | 1.1433
| Median |  |  | 1.2 | 0.585 | 0.9 | 0.75 | 0.65 | 0.61 | 0.875 | 1.32 | 0.75099 | 1.0499 | 1.1 | 0.3 | 0.805 | 0.64 | 0.795 | 0.835 | 1.09
| Std Dev. |  |  | 0.075277 | 0.0098319 | 0.046224 | 0.024221 | 0.018348 | 0.018348 | 0.034303 | 0.028284 | 0.032966 | 0.038474 | 0.040825 | 0.040825 | 0.034303 | 0.048028 | 0.033267 | 0.12822 | 0.4226
| Rel.Std.Dev. |  |  | 6.361% | 1.69% | 5.203% | 3.244% | 2.816% | 3% | 3.935% | 2.143% | 4.417% | 3.615% | 3.768% | 12.89% | 4.244% | 7.544% | 4.176% | 14.91% | 36.96%
| PDM3 |  |  | 49.73% | -26.4% | 12.4% | -5.522% | -17.54% | -22.6% | 10.3% | 67.02% | -5.567% | 34.67% | 37.08% | -59.93% | 2.281% | -19.44% | 0.8051% | 8.819% | 44.67%
| Z-Score (Absolute) |  | 7.28 | NA | 0.98 | 0.34 | 0.27 | 0.67 | 0.85 | 0.27 | 2.2 | 0.27 | 1.1 | NA | NA | 0 | 0.74 | 0.05 | 0.22 | 1.44
Results from laboratories 1.02, 1.17 and 1.18 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Nd in OREAS 627 (Indicative Value 15.7 ppm) | [30 empty] | Indicative
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  |  | 1
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  |  | 1
| 1 | 1 | 1.01: 14.7; 1.02: 17.358; 1.03: 14.3847; 1.04: 17.2924 |  |  | 1
| 1 | 2 | 1.01: 13.9; 1.02: 16.383; 1.03: 14.9183; 1.04: 16.9988 |  |  | 7
| 1 | 3 | 1.01: 14; 1.02: 16.461; 1.03: 15.1287; 1.04: 17.7273 |  |  | 16
| 1 | 4 | 1.01: 14; 1.02: 15.794; 1.03: 15.1716; 1.04: 17.8518 |  |  | 15.6579
| 1 | 5 | 1.01: 13.6; 1.02: 15.929; 1.03: 15.0216; 1.04: 17.862 |  |  | 13
| 1 | 6 | 1.01: 14; 1.02: 15.026; 1.03: 14.7596; 1.04: 17.5217
| Mean |  | 1.01: 14.0333; 1.02: 16.1585; 1.03: 14.8974; 1.04: 17.5423
| Median |  | 1.01: 14; 1.02: 16.156; 1.03: 14.9699; 1.04: 17.6245
| Std Dev. |  | 1.01: 0.36148; 1.02: 0.78087; 1.03: 0.29202; 1.04: 0.34364
| Rel.Std.Dev. |  | 1.01: 2.576%; 1.02: 4.833%; 1.03: 1.96%; 1.04: 1.959%
| PDM3 |  | 1.01: -10.38%; 1.02: 3.197%; 1.03: -4.857%; 1.04: 12.03%
| Z-Score (Absolute) |  | 1.01: 0.95; 1.02: 0.4; 1.03: 0.4; 1.04: 1.28
| Analytical results for Ni in OREAS 627 (Certified Value 7.01 ppm) | [30 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 2
| 1 | 1 | 7 | 6.9 | 5.5 | 7 | 8.2 | 6.2 | 7.2 | 7 | 6.7 | 7.6 | 6.9 | 7 | 6.5472 | 8.71 | 6.1462 | 6.72 | 4 | 5.3 | 15.37 | 7.9 | 6.61 | 7.2 | 11 | 7.4 | 7.1 | 7.3 | 6 | 8.7 |  | 1
| 1 | 2 | 7 | 6.3 | 5.7 | 7 | 7.8 | 6.2 | 6.9 | 7.5 | 7.2 | 7.4 | 6.8 | 6 | 7.402 | 8.76 | 6.3338 | 6.93 | 3 | 5.6 | 15.85 | 7.8 | 6.26 | 7.6 | 12 | 7.3 | 7.5 | 6.8 | 6 | 8.3 |  | 14
| 1 | 3 | 7 | 6.2 | 6.1 | 7.2 | 8.4 | 6 | 6.7 | 7.3 | 7.1 | 7.1 | 6.6 | 6 | 7.0353 | 8.64 | 6.1699 | 6.7 | 4 | 5.7 | 16.13 | 7.8 | 6.48 | 7.8 | 11 | 7.3 | 7.6 | 7 | 6 | 7.9 |  | 16
| 1 | 4 | 7 | 5.9 | 5.7 | 7.3 | 6.6 | 6.3 | 7 | 7.4 | 6.9 | 6.8 | 7.1 | 6 | 7.3645 | 8.708 | 6.0634 | 6.7 | 3 | 6.4 | 15.76 | 7.9 | 6.13 | 7.5 | 12 | 7.3 | 7.8 | 7 | 8 | 8 |  | 7.0092
| 1 | 5 | 7 | 6.1 | 6 | 7.1 | 7.7 | 6.3 | 6.7 | 7.2 | 7.3 | 6.7 | 7.3 | 6 | 7.2454 | 8.39 | 6.3372 | 6.3 | 4 | 5.8 | 15.77 | 8.1 | 6.55 | 7.6 | 12 | 7.1 | 7.8 | 6.9 | 6 | 7.5 |  | 97
| 1 | 6 | 7 | 6.2 | 5.5 | 7.6 | 7.4 | 6.5 | 6.7 | 7.2 | 7.3 | 6.1 | 6.9 | 6 | 6.7702 | 8.264 | 5.9727 | 6.76 | 4 | 5.4 | 15.82 | 8 | 6.23 | 7.1 | 11 | 6.8 | 8.1 | 7 | 7 | 8.3
| Mean |  | 7 | 6.2667 | 5.75 | 7.2 | 7.6833 | 6.25 | 6.8667 | 7.2667 | 7.0833 | 6.95 | 6.9333 | 6.1667 | 7.0608 | 8.5787 | 6.1705 | 6.685 | 3.6667 | 5.7 | 15.7833 | 7.9167 | 6.3767 | 7.4667 | 11.5 | 7.2 | 7.65 | 7 | 6.5 | 8.1167
| Median |  | 7 | 6.2 | 5.7 | 7.15 | 7.75 | 6.25 | 6.8 | 7.25 | 7.15 | 6.95 | 6.9 | 6 | 7.1404 | 8.674 | 6.1581 | 6.71 | 4 | 5.65 | 15.795 | 7.9 | 6.37 | 7.55 | 11.5 | 7.3 | 7.7 | 7 | 6 | 8.15
| Std Dev. |  | 0 | 0.33862 | 0.251 | 0.22804 | 0.64005 | 0.16432 | 0.20656 | 0.17512 | 0.24014 | 0.53944 | 0.24221 | 0.40825 | 0.34397 | 0.2026 | 0.14536 | 0.20763 | 0.5164 | 0.38987 | 0.24394 | 0.1169 | 0.19552 | 0.26583 | 0.54772 | 0.21909 | 0.33912 | 0.16733 | 0.83666 | 0.41191
| Rel.Std.Dev. |  | 0% | 5.404% | 4.365% | 3.167% | 8.33% | 2.629% | 3.008% | 2.41% | 3.39% | 7.762% | 3.493% | 6.62% | 4.872% | 2.362% | 2.356% | 3.106% | 14.08% | 6.84% | 1.546% | 1.477% | 3.066% | 3.56% | 4.763% | 3.043% | 4.433% | 2.39% | 12.87% | 5.075%
| PDM3 |  | -0.131% | -10.59% | -17.96% | 2.722% | 9.618% | -10.83% | -2.033% | 3.674% | 1.058% | -0.8443% | -1.082% | -12.02% | 0.736% | 22.39% | -11.96% | -4.625% | -47.69% | -18.68% | 125.2% | 12.95% | -9.024% | 6.527% | 64.07% | 2.722% | 9.143% | -0.131% | -7.264% | 15.8%
| Z-Score (Absolute) |  | NA | 0.91 | 1.5 | 0.16 | 0.71 | 0.93 | 0.22 | 0.24 | 0.03 | 0.13 | 0.15 | NA | 0 | 1.74 | 1.02 | 0.43 | NA | 1.56 | 9.98 | 0.98 | 0.78 | 0.46 | NA | 0.16 | 0.67 | 0.07 | NA | 1.21
Results from laboratories 1.01, 1.12, 1.17, 1.24 and 1.28 were removed due to their 1 ppm reading resolution.
| Analytical results for P in OREAS 627 (Certified Value 0.032 wt.%) | [30 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 3
| 1 | 1 | 0.0321 | 0.036 | 0.029 | 0.031 | 0.035 | 0.032 | 0.032 | 0.032 | 0.0302 | 0.0333 | 0.03 | 0.0324 | 0.030653 | 0.034596 | 0.03 | 0.0331 | 0.026 | 0.0341 | 0.03 | 0.03 | 0.028 | 0.03 | 0.0331 | 0.0321 | 0.032 | 0.0316 |  |  |  | 1
| 1 | 2 | 0.0298 | 0.034 | 0.03 | 0.032 | 0.033 | 0.032 | 0.031 | 0.033 | 0.0303 | 0.0322 | 0.03 | 0.0308 | 0.032857 | 0.033151 | 0.03 | 0.032 | 0.026 | 0.0345 | 0.03 | 0.03 | 0.029 | 0.03 | 0.0323 | 0.0314 | 0.033 | 0.0314 |  |  |  | 24
| 1 | 3 | 0.0334 | 0.034 | 0.031 | 0.031 | 0.037 | 0.032 | 0.03 | 0.032 | 0.0317 | 0.0312 | 0.028 | 0.0314 | 0.032089 | 0.034369 | 0.03 | 0.033 | 0.027 | 0.0347 | 0.03 | 0.03 | 0.029 | 0.03 | 0.0324 | 0.032 | 0.033 | 0.0311 |  |  |  | 16
| 1 | 4 | 0.031 | 0.035 | 0.031 | 0.031 | 0.032 | 0.035 | 0.031 | 0.033 | 0.0305 | 0.0303 | 0.03 | 0.0307 | 0.032224 | 0.033134 | 0.03 | 0.0326 | 0.027 | 0.0344 | 0.03 | 0.03 | 0.029 | 0.03 | 0.0335 | 0.0327 | 0.034 | 0.0303 |  |  |  | 0.031592
| 1 | 5 | 0.0307 | 0.034 | 0.031 | 0.032 | 0.033 | 0.032 | 0.031 | 0.032 | 0.0311 | 0.0305 | 0.031 | 0.0314 | 0.032111 | 0.032688 | 0.03 | 0.032 | 0.027 | 0.0344 | 0.03 | 0.03 | 0.028 | 0.03 | 0.033 | 0.0307 | 0.034 | 0.0299 |  |  |  | 98
| 1 | 6 | 0.0333 | 0.035 | 0.03 | 0.033 | 0.033 | 0.033 | 0.031 | 0.031 | 0.0309 | 0.0303 | 0.03 | 0.0312 | 0.031262 | 0.033096 | 0.03 | 0.0324 | 0.026 | 0.0348 | 0.03 | 0.03 | 0.029 | 0.03 | 0.0334 | 0.032 | 0.032 | 0.0305
| Mean |  | 0.031717 | 0.034667 | 0.030333 | 0.031667 | 0.033833 | 0.032667 | 0.031 | 0.032167 | 0.030783 | 0.0313 | 0.029833 | 0.031317 | 0.031866 | 0.033506 | 0.03 | 0.032517 | 0.0265 | 0.034483 | 0.03 | 0.03 | 0.028667 | 0.03 | 0.03295 | 0.031817 | 0.033 | 0.0308
| Median |  | 0.03155 | 0.0345 | 0.0305 | 0.0315 | 0.033 | 0.032 | 0.031 | 0.032 | 0.0307 | 0.03085 | 0.03 | 0.0313 | 0.0321 | 0.033143 | 0.03 | 0.0325 | 0.0265 | 0.03445 | 0.03 | 0.03 | 0.029 | 0.03 | 0.03305 | 0.032 | 0.033 | 0.0308
| Std Dev. |  | 0.0014634 | 0.0008165 | 0.0008165 | 0.0008165 | 0.0018348 | 0.0012111 | 0.00063246 | 0.00075277 | 0.00056716 | 0.0012215 | 0.00098319 | 0.000608 | 0.00078186 | 0.00077905 | 0 | 0.00047504 | 0.00054772 | 0.00024833 | 0 | 0 | 0.0005164 | 0 | 0.000501 | 0.00068532 | 0.00089443 | 0.00066933
| Rel.Std.Dev. |  | 4.614% | 2.355% | 2.692% | 2.578% | 5.423% | 3.707% | 2.04% | 2.34% | 1.842% | 3.902% | 3.296% | 1.941% | 2.454% | 2.325% | 0% | 1.461% | 2.067% | 0.7201% | 0% | 0% | 1.801% | 0% | 1.52% | 2.154% | 2.71% | 2.173%
| PDM3 |  | 0.3961% | 9.734% | -3.983% | 0.2378% | 7.096% | 3.403% | -1.872% | 1.821% | -2.558% | -0.9228% | -5.565% | -0.8701% | 0.8691% | 6.059% | -5.038% | 2.928% | -16.12% | 9.154% | -5.038% | -5.038% | -9.258% | -5.038% | 4.3% | 0.7126% | 4.458% | -2.506%
| Z-Score (Absolute) |  | 0.12 | 1.63 | 0.59 | 0.09 | 1.2 | 0.6 | 0.25 | 0.35 | 0.36 | 0.1 | 0.85 | 0.09 | 0.19 | 1.03 | 0.76 | 0.52 | 2.56 | 1.53 | 0.76 | 0.76 | 1.45 | 0.76 | 0.75 | 0.17 | 0.77 | 0.35
| Analytical results for Pb in OREAS 627 (Certified Value 2.24 wt.%) | [30 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*AAS | AR*OES | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*AAS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*AAS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*OES | AR*OES | AR*OES/MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.2g | 0.5g | 0.1g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 0.25g | 01g | 0.25g | 0.25g | 0.4g | 0.25g | 0.2g |  | 3
| 1 | 1 | 2.28 | 2.14 | 2.069 | 2.24 | 2.16 | 2.22 | 2.22 | 2.26 | 2.38 | > 0.5 | 2.19 | > 0.5 | 2.2284 | 2.0603 | 2.3154 | 2.28 | > 0.5 | > 1 | 2.2214 | 2.39 | 2.261 | 2.252 | 2.23 | > 1 | 2.25 | 2.197 | > 1 | 2.3145 |  | 1
| 1 | 2 | 2.3 | 2.12 | 2.104 | 2.23 | 2.09 | 2.22 | 2.23 | 2.26 | 2.42 | > 0.5 | 2.18 | > 0.5 | 2.2435 | 2.425 | 2.3193 | 2.26 | > 0.5 | > 1 | 2.2356 | 2.3864 | 2.251 | 2.24 | 2.2 | > 1 | 2.21 | 2.203 | > 1 | 2.2914 |  | 25
| 1 | 3 | 2.21 | 2.16 | 2.111 | 2.25 | 2.17 | 2.2 | 2.24 | 2.28 | 2.37 | > 0.5 | 2.22 | > 0.5 | 2.2523 | 2.5039 | 2.3043 | 2.27 | > 0.5 | > 1 | 2.2563 | 2.3845 | 2.28 | 2.291 | 2.17 | > 1 | 2.21 | 2.214 | > 1 | 2.2746 |  | 16
| 1 | 4 | 2.27 | 2.16 | 2.186 | 2.25 | 2.14 | 2.23 | 2.22 | 2.41 | 2.4 | > 0.5 | 2.19 | > 0.5 | 2.2611 | 2.4797 | 2.3834 | 2.27 | > 0.5 | > 1 | 2.2347 | 2.2765 | 2.285 | 2.212 | 2.23 | > 1 | 2.22 | 2.222 | > 1 | 2.2888 |  | 2.2375
| 1 | 5 | 2.33 | 2.18 | 2.125 | 2.22 | 2.15 | 2.23 | 2.23 | 2.27 | 2.41 | > 0.5 | 2.18 | > 0.5 | 2.2646 | 2.2331 | 2.3718 | 2.27 | > 0.5 | > 1 | 2.2377 | 2.345 | 2.294 | 2.281 | 2.22 | > 1 | 2.24 | 2.235 | > 1 | 2.2582 |  | 99
| 1 | 6 | 2.35 | 2.18 | 2.142 | 2.23 | 2.16 | 2.24 | 2.21 | 2.29 | 2.39 | > 0.5 | 2.19 | > 0.5 | 2.2738 | 2.0419 | 2.3252 | 2.28 | > 0.5 | > 1 | 2.2359 | 2.3306 | 2.281 | 2.279 | 2.24 | > 1 | 2.22 | 2.187 | > 1 | 2.3045
| Mean |  | 2.29 | 2.1567 | 2.1228 | 2.2367 | 2.145 | 2.2233 | 2.225 | 2.295 | 2.395 |  | 2.1917 |  | 2.2539 | 2.2906 | 2.3366 | 2.2717 |  |  | 2.2369 | 2.3522 | 2.2753 | 2.2592 | 2.215 |  | 2.225 | 2.2097 |  | 2.2887
| Median |  | 2.29 | 2.16 | 2.118 | 2.235 | 2.155 | 2.225 | 2.225 | 2.275 | 2.395 |  | 2.19 |  | 2.2567 | 2.329 | 2.3223 | 2.27 |  |  | 2.2357 | 2.3647 | 2.2805 | 2.2655 | 2.225 |  | 2.22 | 2.2085 |  | 2.2901
| Std Dev. |  | 0.049396 | 0.023381 | 0.039382 | 0.012111 | 0.02881 | 0.013663 | 0.010488 | 0.057533 | 0.018708 |  | 0.01472 |  | 0.016273 | 0.20858 | 0.032741 | 0.0075277 |  |  | 0.011178 | 0.044476 | 0.016083 | 0.030089 | 0.025884 |  | 0.016432 | 0.0175 |  | 0.020257
| Rel.Std.Dev. |  | 2.157% | 1.084% | 1.855% | 0.5415% | 1.343% | 0.6145% | 0.4714% | 2.507% | 0.7811% |  | 0.6716% |  | 0.722% | 9.106% | 1.401% | 0.3314% |  |  | 0.4997% | 1.891% | 0.7068% | 1.332% | 1.169% |  | 0.7385% | 0.792% |  | 0.8851%
| PDM3 |  | 2.345% | -3.614% | -5.126% | -0.03904% | -4.136% | -0.6349% | -0.5604% | 2.568% | 7.037% |  | -2.05% |  | 0.7332% | 2.373% | 4.425% | 1.525% |  |  | -0.02712% | 5.123% | 1.689% | 0.9665% | -1.007% |  | -0.5604% | -1.246% |  | 2.285%
| Z-Score (Absolute) |  | 0.76 | 1.52 | 2.09 | 0.15 | 1.71 | 0.38 | 0.35 | 0.85 | 2.55 | NA | 0.92 | NA | 0.15 | 0.77 | 1.56 | 0.45 | NA | NA | 0.15 | 1.82 | 0.51 | 0.23 | 0.52 | NA | 0.35 | 0.61 | NA | 0.74
| Analytical results for Pd in OREAS 627 (Indicative Value 13.2 ppb) | [30 empty] | Indicative
Pd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.15 |  | ppb
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*OES/MS |  | 1
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 01g |  | 1
| 1 | 1 | 1.01: 11; 1.02: < 10; 1.03: < 10 |  | 1
| 1 | 2 | 1.01: 15; 1.02: < 10; 1.03: < 10 |  | 1
| 1 | 3 | 1.01: 11; 1.02: < 10; 1.03: < 10 |  | 16
| 1 | 4 | 1.01: 11; 1.02: < 10; 1.03: < 10 |  | 13.1667
| 1 | 5 | 1.01: 14; 1.02: < 10; 1.03: < 10 |  | 14
| 1 | 6 | 1.01: 17; 1.02: < 10; 1.03: < 10
| Mean |  | 1.01: 13.1667
| Median |  | 1.01: 12.5
| Std Dev. |  | 1.01: 2.5626
| Rel.Std.Dev. |  | 1.01: 19.46%
| PDM3 |  | 1.01: -2.554e-13%
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Pr in OREAS 627 (Indicative Value 3.99 ppm) | [30 empty] | Indicative
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g |  | 2
| 1 | 1 | 1.01: 3.6; 1.02: 4.388; 1.03: 3.8333; 1.04: 4.6287 |  | 1
| 1 | 2 | 1.01: 3.3; 1.02: 4.27; 1.03: 3.9836; 1.04: 4.5674 |  | 9
| 1 | 3 | 1.01: 3.4; 1.02: 4.184; 1.03: 4.0253; 1.04: 4.5904 |  | 16
| 1 | 4 | 1.01: 3.4; 1.02: 4.092; 1.03: 4.025; 1.04: 4.2785 |  | 3.9915
| 1 | 5 | 1.01: 3.3; 1.02: 4.125; 1.03: 3.9567; 1.04: 4.3172 |  | 15
| 1 | 6 | 1.01: 3.3; 1.02: 3.891; 1.03: 3.9199; 1.04: 4.4192
| Mean |  | 1.01: 3.3833; 1.02: 4.1583; 1.03: 3.9573; 1.04: 4.4669
| Median |  | 1.01: 3.35; 1.02: 4.1545; 1.03: 3.9702; 1.04: 4.4933
| Std Dev. |  | 1.01: 0.1169; 1.02: 0.16903; 1.03: 0.073008; 1.04: 0.1495
| Rel.Std.Dev. |  | 1.01: 3.455%; 1.02: 4.065%; 1.03: 1.845%; 1.04: 3.347%
| PDM3 |  | 1.01: -15.24%; 1.02: 4.181%; 1.03: -0.8561%; 1.04: 11.91%
| Z-Score (Absolute) |  | 1.01: 1.78; 1.02: 0.27; 1.03: 0.27; 1.04: 1.08
| Analytical results for Pt in OREAS 627 (Indicative Value < 5 ppb) | [30 empty] | Indicative
Pt | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.15 |  | ppb
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*OES/MS |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 01g |  | 2
| 1 | 1 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 1
| 1 | 2 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 10
| 1 | 3 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 16
| 1 | 4 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | < 5
| 1 | 5 | 1.01: < 10; 1.02: < 5; 1.03: < 5 |  | 16
| 1 | 6 | 1.01: < 10; 1.02: < 5; 1.03: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for Rb in OREAS 627 (Certified Value 17.7 ppm) | [30 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [12 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [12 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [12 empty] | 1
| 1 | 1 | 18.4 | 17 | 21.7 | 14.4 | 17.6 | 16.8 | 14.3 | 20.67 | 16.967 | 19.418 | 23.7 | 14.3 | 20 | 17 | 17.4 | 17.9 | 19.8 | [12 empty] | 1
| 1 | 2 | 17.9 | 17.1 | 19.5 | 15 | 18 | 17.3 | 14.3 | 20.45 | 17.3758 | 19.2844 | 23.5 | 14.7 | 20.2 | 17.9 | 17.7 | 17.5 | 18.3 | [12 empty] | 43
| 1 | 3 | 19.1 | 17 | 22.3 | 15.2 | 17.2 | 17.1 | 14.7 | 19.91 | 17.6114 | 18.8172 | 23.9 | 14.6 | 19.3 | 18.5 | 17.9 | 17.1 | 19.3 | [12 empty] | 16
| 1 | 4 | 17.9 | 16.8 | 19 | 14.4 | 17.8 | 17 | 14.3 | 19.95 | 17.4948 | 18.2202 | 23.6 | 14.5 | 19.6 | 17.3 | 17.9 | 17.5 | 19.2 | [12 empty] | 17.6513
| 1 | 5 | 17.7 | 17.2 | 21 | 15.1 | 17.4 | 16.8 | 14.8 | 19.95 | 17.4151 | 18.4016 | 23.3 | 14.7 | 20.3 | 17.5 | 18.1 | 17.6 | 18.7 | [12 empty] | 100
| 1 | 6 | 18.5 | 17.5 | 19.9 | 15 | 17.4 | 16.6 | 14.6 | 19 | 16.8825 | 18.704 | 23.9 | 14.3 | 19.4 | 17.5 | 18.4 | 18.5 | 19
| Mean |  | 18.25 | 17.1 | 20.5667 | 14.85 | 17.5667 | 16.9333 | 14.5 | 19.9883 | 17.2911 | 18.8076 | 23.65 | 14.5167 | 19.8 | 17.6167 | 17.9 | 17.6833 | 19.05
| Median |  | 18.15 | 17.05 | 20.45 | 15 | 17.5 | 16.9 | 14.45 | 19.95 | 17.3954 | 18.7606 | 23.65 | 14.55 | 19.8 | 17.5 | 17.9 | 17.55 | 19.1
| Std Dev. |  | 0.52058 | 0.23664 | 1.3049 | 0.35637 | 0.29439 | 0.25033 | 0.22804 | 0.57683 | 0.2962 | 0.47345 | 0.23452 | 0.18348 | 0.42426 | 0.52313 | 0.34059 | 0.47504 | 0.51672
| Rel.Std.Dev. |  | 2.852% | 1.384% | 6.345% | 2.4% | 1.676% | 1.478% | 1.573% | 2.886% | 1.713% | 2.517% | 0.9916% | 1.264% | 2.143% | 2.97% | 1.903% | 2.686% | 2.712%
| PDM3 |  | 3.392% | -3.123% | 16.52% | -15.87% | -0.4793% | -4.067% | -17.85% | 13.24% | -2.04% | 6.551% | 33.98% | -17.76% | 12.17% | -0.196% | 1.409% | 0.1816% | 7.924%
| Z-Score (Absolute) |  | 0.34 | 0.35 | 1.73 | 1.7 | 0.07 | 0.45 | 1.91 | 1.38 | 0.24 | 0.67 | 3.58 | 1.9 | 1.27 | 0.04 | 0.13 | 0 | 0.82
| Analytical results for Re in OREAS 627 (Indicative Value 0.001 ppm) | [30 empty] | Indicative
Re | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab; 1.14: Lab; 1.15: Lab; 1.16: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.04; 1.03: 1.05; 1.04: 1.06; 1.05: 1.07; 1.06: 1.08; 1.07: 1.09; 1.08: 1.10; 1.09: 1.13; 1.10: 1.15; 1.11: 1.17; 1.12: 1.18; 1.13: 1.25; 1.14: 1.26; 1.15: 1.27; 1.16: 1.30 |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*MS; 1.09: AR*MS; 1.10: AR*OES/MS; 1.11: AR*OES/MS; 1.12: AR*OES/MS; 1.13: AR*OES/MS; 1.14: AR*OES/MS; 1.15: AR*OES/MS; 1.16: AR*MS |  | 3
|  |  | 1.01: 0.5g; 1.02: 0.5g; 1.03: 0.5g; 1.04: 0.5g; 1.05: 0.5g; 1.06: 0.5g; 1.07: 0.5g; 1.08: 25g; 1.09: 0.2g; 1.10: 01g; 1.11: 01g; 1.12: 0.25g; 1.13: 0.25g; 1.14: 0.25g; 1.15: 0.5g; 1.16: 0.2g |  | 3
| 1 | 1 | 1.01: < 0.001; 1.02: 0.001; 1.03: 0.001; 1.04: 0.001; 1.05: 0.002; 1.06: 0.001; 1.07: 0.001; 1.08: < 0.001; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.002; 1.15: 0.001; 1.16: 0.003 |  | 1
| 1 | 2 | 1.01: < 0.001; 1.02: 0.001; 1.03: 0.001; 1.04: 0.002; 1.05: 0.001; 1.06: 0.001; 1.07: 0.002; 1.08: 0.002; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.002; 1.15: 0.001; 1.16: 0.002 |  | 15
| 1 | 3 | 1.01: < 0.001; 1.02: 0.001; 1.03: 0.001; 1.04: 0.001; 1.05: 0.001; 1.06: 0.001; 1.07: 0.002; 1.08: < 0.001; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.002; 1.15: 0.001; 1.16: < 0.002 |  | 16
| 1 | 4 | 1.01: < 0.001; 1.02: 0.001; 1.03: 0.001; 1.04: 0.001; 1.05: 0.001; 1.06: 0.001; 1.07: 0.002; 1.08: < 0.001; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.003; 1.15: 0.001; 1.16: 0.003 |  | 0.0012917
| 1 | 5 | 1.01: < 0.001; 1.02: 0.002; 1.03: 0.001; 1.04: 0.002; 1.05: 0.001; 1.06: 0.002; 1.07: 0.001; 1.08: 0.001; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.002; 1.15: 0.002; 1.16: 0.002 |  | 12
| 1 | 6 | 1.01: < 0.001; 1.02: 0.002; 1.03: 0.001; 1.04: 0.001; 1.05: 0.001; 1.06: 0.001; 1.07: 0.002; 1.08: < 0.001; 1.09: < 0.01; 1.10: < 0.05; 1.11: < 0.05; 1.12: < 0.005; 1.13: < 0.02; 1.14: 0.002; 1.15: 0.001; 1.16: 0.004
| Mean |  | 1.02: 0.0013333; 1.03: 0.001; 1.04: 0.0013333; 1.05: 0.0011667; 1.06: 0.0011667; 1.07: 0.0016667; 1.08: 0.0015; 1.14: 0.0021667; 1.15: 0.0011667; 1.16: 0.0028
| Median |  | 1.02: 0.001; 1.03: 0.001; 1.04: 0.001; 1.05: 0.001; 1.06: 0.001; 1.07: 0.002; 1.08: 0.0015; 1.14: 0.002; 1.15: 0.001; 1.16: 0.003
| Std Dev. |  | 1.02: 0.0005164; 1.03: 0; 1.04: 0.0005164; 1.05: 0.00040825; 1.06: 0.00040825; 1.07: 0.0005164; 1.08: 0.00070711; 1.14: 0.00040825; 1.15: 0.00040825; 1.16: 0.00083666
| Rel.Std.Dev. |  | 1.02: 38.73%; 1.03: 0%; 1.04: 38.73%; 1.05: 34.99%; 1.06: 34.99%; 1.07: 30.98%; 1.08: 47.14%; 1.14: 18.84%; 1.15: 34.99%; 1.16: 29.88%
| PDM3 |  | 1.02: 3.226%; 1.03: -22.58%; 1.04: 3.226%; 1.05: -9.677%; 1.06: -9.677%; 1.07: 29.03%; 1.08: 16.13%; 1.14: 67.74%; 1.15: -9.677%; 1.16: 116.8%
| Z-Score (Absolute) |  | 1.01: 1.01; 1.02: 0.17; 1.03: 0.51; 1.04: 0.17; 1.05: 0.34; 1.06: 0.34; 1.07: 0.17; 1.08: 0.67; 1.09: 3.54; 1.10: 23.77; 1.11: 23.77; 1.12: 1.01; 1.13: 8.6; 1.14: 0.67; 1.15: 0.34; 1.16: 1.01
| Analytical results for S in OREAS 627 (Certified Value 9.71 wt.%) | [30 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  | 3
| 1 | 1 | 9.92 | 10.54 | 9.99 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.45 | > 5 | 9.45 | > 5 | 9.9086 | > 5 | 8.38 | > 5 | 9.62 | 9.7257 | > 10 | > 5 | 9.65 | > 5 | > 5 | > 10 | > 5 | 6.75 |  |  |  | 1
| 1 | 2 | 9.55 | 9.993 | 9.97 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.52 | > 5 | 9.66 | > 5 | 10.0325 | > 5 | 8.31 | > 5 | 9.17 | 9.6545 | > 10 | > 5 | 9.54 | > 5 | > 5 | > 10 | > 5 | 6.98 |  |  |  | 26
| 1 | 3 | 9.71 | 10.22 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.93 | > 5 | 9.2 | > 5 | 10.0097 | > 5 | 8.25 | > 5 | 9.58 | 9.6422 | > 10 | > 5 | 9.43 | > 5 | > 5 | > 10 | > 5 | 7.14 |  |  |  | 16
| 1 | 4 | 9.7 | 10.42 | > 10 | > 10 | > 10 | > 10 | 9.94 | > 10 | 8.57 | > 5 | 9.35 | > 5 | 10.2116 | > 5 | 8.21 | > 5 | 9.64 | 9.8507 | > 10 | > 5 | 9.61 | > 5 | > 5 | > 10 | > 5 | 6.85 |  |  |  | 9.7098
| 1 | 5 | 9.76 | 10.01 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.73 | > 5 | 9.65 | > 5 | 10.1429 | > 5 | 8.72 | > 5 | 9.74 | 9.7475 | > 10 | > 5 | 9.39 | > 5 | > 5 | 9.99 | > 5 | 7.84 |  |  |  | 101
| 1 | 6 | 9.45 | 10.4 | > 10 | > 10 | > 10 | > 10 | > 10 | > 10 | 8.45 | > 5 | 9.07 | > 5 | 10.1287 | > 5 | 8.31 | > 5 | 9.64 | 9.5476 | > 10 | > 5 | 9.6 | > 5 | > 5 | > 10 | > 5 | 6.49
| Mean |  | 1.01: 9.6817; 1.02: 10.2638; 1.03: 9.98; 1.07: 9.94; 1.09: 8.6083; 1.11: 9.3967; 1.13: 10.0723; 1.16: 8.3633; 1.18: 9.565; 1.20: 9.6947; 1.24: 9.5367; 1.27: 9.99; 1.30: 7.0083
| Median |  | 1.01: 9.705; 1.02: 10.31; 1.03: 9.98; 1.07: 9.94; 1.09: 8.545; 1.11: 9.4; 1.13: 10.0806; 1.16: 8.31; 1.18: 9.63; 1.20: 9.6901; 1.24: 9.57; 1.27: 9.99; 1.30: 6.915
| Std Dev. |  | 1.01: 0.16437; 1.02: 0.22755; 1.03: 0.014142; 1.09: 0.18851; 1.11: 0.2383; 1.13: 0.10942; 1.16: 0.18414; 1.18: 0.20057; 1.20: 0.10401; 1.24: 0.10501; 1.30: 0.46266
| Rel.Std.Dev. |  | 1.01: 1.698%; 1.02: 2.217%; 1.03: 0.1417%; 1.09: 2.19%; 1.11: 2.536%; 1.13: 1.086%; 1.16: 2.202%; 1.18: 2.097%; 1.20: 1.073%; 1.24: 1.101%; 1.30: 6.602%
| PDM3 |  | 1.01: -0.2901%; 1.02: 5.706%; 1.03: 2.782%; 1.07: 2.37%; 1.09: -11.34%; 1.11: -3.225%; 1.13: 3.733%; 1.16: -13.87%; 1.18: -1.492%; 1.20: -0.1559%; 1.24: -1.783%; 1.27: 2.885%; 1.30: -27.82%
| Z-Score (Absolute) |  | 0 | 1.32 | 0.67 | NA | NA | NA | 0.58 | NA | 2.43 | NA | 0.64 | NA | 0.88 | NA | 2.98 | NA | 0.26 | 0.03 | NA | NA | 0.33 | NA | NA | 0.7 | NA | 6.04
| Analytical results for Sb in OREAS 627 (Certified Value 113 ppm) | [30 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  | 0
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  | 0
| 1 | 1 | 108 | 107 | 127.3 | 112 | 128 | 124 | 119.5 | 124 | 123 | 121.1 | 102.05 | 103 | 105.9272 | 82.677 | 101.6173 | 129.8 | 113 | 91.46 | 108 | 118.07 | 120 | 122 | 113 | 108 | 86.85 | 108.3 | 114 | 125.49 |  | 1
| 1 | 2 | 105 | 108 | 126.7 | 111 | 118 | 130 | 117.5 | 124.5 | 124 | 117.01 | 104.49 | 102 | 107.8183 | 98.49 | 100.6566 | 125.4 | 101 | 90.68 | 111 | 117.24 | 117 | 122 | 116 | 113 | 90.64 | 106.5 | 114 | 112.31 |  | 45
| 1 | 3 | 109 | 104 | 129.8 | 110.5 | 132 | 129.5 | 116 | 124 | 127 | 117.36 | 97.89 | 106 | 110.074 | 100.858 | 101.9036 | 127.6 | 106 | 88.47 | 109 | 114.22 | 117 | 120 | 108 | 115 | 91.24 | 109.77 | 129 | 114.86 |  | 16
| 1 | 4 | 107 | 112 | 129 | 110.5 | 122.5 | 127.5 | 115 | 126 | 126 | 111.68 | 99.72 | 100 | 108.8761 | 96.855 | 101.9337 | 129.6 | 107 | 89.07 | 114 | 115.15 | 117 | 117 | 111 | 114 | 89.27 | 109 | 128 | 122.09 |  | 112.554
| 1 | 5 | 109 | 103 | 128.6 | 111.5 | 125.5 | 123.5 | 115 | 124 | 128 | 115.45 | 103.34 | 103 | 108.9219 | 99.612 | 101.6268 | 126.1 | 111 | 92.4 | 107 | 117.69 | 119 | 121 | 111 | 113 | 90 | 108.02 | 118 | 125.24 |  | 102
| 1 | 6 | 108 | 103 | 127.3 | 110 | 122.5 | 128.5 | 117.5 | 124 | 125 | 111.78 | 92.68 | 103 | 107.8789 | 99.38 | 100.8664 | 134.1 | 107 | 92.87 | 106 | 119.32 | 115 | 118 | 110 | 109 | 91.51 | 107.42 | 116 | 123.74
| Mean |  | 107.6667 | 106.1667 | 128.1167 | 110.9167 | 124.75 | 127.1667 | 116.75 | 124.4167 | 125.5 | 115.73 | 100.0283 | 102.8333 | 108.2494 | 96.312 | 101.434 | 128.7667 | 107.5 | 90.825 | 109.1667 | 116.9483 | 117.5 | 120 | 111.5 | 112 | 89.9183 | 108.1683 | 119.8333 | 120.6217
| Median |  | 108 | 105.5 | 127.95 | 110.75 | 124 | 128 | 116.75 | 124 | 125.5 | 116.23 | 100.885 | 103 | 108.3775 | 98.935 | 101.622 | 128.6 | 107 | 91.07 | 108.5 | 117.465 | 117 | 120.5 | 111 | 113 | 90.32 | 108.16 | 117 | 122.915
| Std Dev. |  | 1.5055 | 3.5449 | 1.199 | 0.73598 | 4.8862 | 2.7869 | 1.7536 | 0.80104 | 1.8708 | 3.6115 | 4.3273 | 1.9408 | 1.4064 | 6.8106 | 0.54177 | 3.1627 | 4.1833 | 1.7727 | 2.9269 | 1.9077 | 1.7607 | 2.0976 | 2.7386 | 2.8284 | 1.7113 | 1.1524 | 6.8823 | 5.6425
| Rel.Std.Dev. |  | 1.398% | 3.339% | 0.9359% | 0.6635% | 3.917% | 2.192% | 1.502% | 0.6438% | 1.491% | 3.121% | 4.326% | 1.887% | 1.299% | 7.071% | 0.5341% | 2.456% | 3.891% | 1.952% | 2.681% | 1.631% | 1.498% | 1.748% | 2.456% | 2.525% | 1.903% | 1.065% | 5.743% | 4.678%
| PDM3 |  | -4.342% | -5.675% | 13.83% | -1.455% | 10.84% | 12.98% | 3.728% | 10.54% | 11.5% | 2.822% | -11.13% | -8.636% | -3.824% | -14.43% | -9.88% | 14.4% | -4.49% | -19.31% | -3.01% | 3.904% | 4.394% | 6.615% | -0.9365% | -0.4922% | -20.11% | -3.897% | 6.467% | 7.168%
| Z-Score (Absolute) |  | 0.34 | 0.46 | 1.35 | 0.07 | 1.07 | 1.27 | 0.41 | 1.05 | 1.14 | 0.33 | 0.97 | 0.74 | 0.29 | 1.27 | 0.85 | 1.41 | 0.35 | 1.73 | 0.21 | 0.43 | 0.47 | 0.68 | 0.02 | 0.02 | 1.8 | 0.3 | 0.67 | 0.73
| Analytical results for Sc in OREAS 627 (Certified Value 1.31 ppm) | [30 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  |  | 2
| 1 | 1 | < 5 | 1.3 | 1.2 | 1.3 | 2 | 1.1 | 1.3 | 1.6 | 1 | 1.2 | 1.1 | 2 | 1.1187 | 1.3286 | 1 | 1 | 0.5 | 2 | 1.5 | < 5 | 1.9 | 1 | 1.4 | 1.3 | 1.5 |  |  |  |  | 1
| 1 | 2 | < 5 | 1.3 | 1.2 | 1.3 | 1.8 | 1.2 | 1.2 | 1.6 | 1 | 1.1 | 1.1 | 1 | 1.1635 | 1.3176 | < 1 | 1 | 0.5 | 1.9 | 1.5 | < 5 | 2 | 1 | 1.4 | 1.3 | 1.3 |  |  |  |  | 46
| 1 | 3 | < 5 | 1.5 | 1.2 | 1.3 | 2 | 1.2 | 1.2 | 1.6 | 1 | 1.1 | 1 | 2 | 1.1284 | 1.2808 | 1 | 1 | 0.5 | 1.8 | 1.5 | < 5 | 2 | 1 | 1.4 | 1.3 | 1.4 |  |  |  |  | 16
| 1 | 4 | < 5 | 1.3 | 1.3 | 1.3 | 1.6 | 1.1 | 1.2 | 1.6 | 1 | 1 | 1 | 2 | 1.1588 | 1.3032 | < 1 | 1 | 0.5 | 1.9 | 1.5 | < 5 | 2.1 | 1 | 1.4 | 1.3 | 1.4 |  |  |  |  | 1.3064
| 1 | 5 | < 5 | 1.3 | 1.3 | 1.3 | 1.9 | 1.2 | 1.2 | 1.6 | 0.9 | 1 | 1.1 | 2 | 1.1379 | 1.3413 | 1.1 | 1 | 0.5 | 2 | 1.5 | < 5 | 2.1 | 1 | 1.4 | 1.3 | 1.4 |  |  |  |  | 103
| 1 | 6 | < 5 | 1.4 | 1.2 | 1.3 | 1.9 | 1.2 | 1.2 | 1.6 | 1 | 0.9 | 1 | 2 | 1.1145 | 1.2914 | 1 | 1 | 0.5 | 1.9 | 1.4 | < 5 | 2.2 | 1.1 | 1.5 | 1.3 | 1.4
| Mean |  |  | 1.35 | 1.2333 | 1.3 | 1.8667 | 1.1667 | 1.2167 | 1.6 | 0.98333 | 1.05 | 1.05 | 1.8333 | 1.137 | 1.3105 | 1.025 | 1 | 0.5 | 1.9167 | 1.4833 |  | 2.05 | 1.0167 | 1.4167 | 1.3 | 1.4
| Median |  |  | 1.3 | 1.2 | 1.3 | 1.9 | 1.2 | 1.2 | 1.6 | 1 | 1.05 | 1.05 | 2 | 1.1331 | 1.3104 | 1 | 1 | 0.5 | 1.9 | 1.5 |  | 2.05 | 1 | 1.4 | 1.3 | 1.4
| Std Dev. |  |  | 0.083666 | 0.05164 | 0 | 0.15055 | 0.05164 | 0.040825 | 2.4324e-16 | 0.040825 | 0.10488 | 0.054772 | 0.40825 | 0.020472 | 0.022912 | 0.05 | 0 | 0 | 0.075277 | 0.040825 |  | 0.10488 | 0.040825 | 0.040825 | 0 | 0.063246
| Rel.Std.Dev. |  |  | 6.197% | 4.187% | 0% | 8.065% | 4.426% | 3.355% | 1.52e-14% | 4.152% | 9.989% | 5.216% | 22.27% | 1.801% | 1.748% | 4.878% | 0% | 0% | 3.928% | 2.752% |  | 5.116% | 4.016% | 2.882% | 0% | 4.518%
| PDM3 |  |  | 3.334% | -5.596% | -0.4933% | 42.88% | -10.7% | -6.872% | 22.47% | -24.73% | -19.63% | -19.63% | 40.33% | -12.97% | 0.3096% | -21.54% | -23.46% | -61.73% | 46.71% | 13.54% |  | 56.91% | -22.18% | 8.437% | -0.4933% | 7.161%
| Z-Score (Absolute) |  | 3.24 | 0.13 | 0.18 | 0 | 1.53 | 0.36 | 0.22 | 0.81 | 0.85 | 0.67 | 0.67 | NA | 0.44 | 0.03 | 1.21 | NA | 2.16 | 1.66 | 0.49 | 3.24 | 2.02 | 0.76 | 0.31 | 0 | 0.27
| Analytical results for Se in OREAS 627 (Indicative Value 2.36 ppm) | [30 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 2
| 1 | 1 | < 5 | 4.3 | 2 | 7.1 | 4.8 | 3.7 | 3.1 | 10.5 | 0.7 | 1 | 1.9 | < 1.5 | 1.2865 | 1 | 1.3 | < 10 | 0.68 | 1.8 | 3 | 5 | 5 | 0.7 | < 10 | [6 empty] | 1
| 1 | 2 | < 5 | 4.3 | 2 | 7 | 4.3 | 2.4 | 2.9 | 11.2 | 0.7 | 1 | 1.9 | < 1.5 | 1.3934 | 1 | 1.2 | < 10 | 0.67 | 1.4 | 3 | 6 | 6 | 0.6 | < 10 | [6 empty] | 16
| 1 | 3 | < 5 | 4.1 | 2 | 6.7 | 4.4 | 2.4 | 3.4 | 10.2 | 0.7 | 1 | 1.8 | < 1.5 | 1.6656 | 1 | 1.4 | < 10 | 0.64 | 1.4 | 2 | 5 | 6 | 0.6 | < 10 | [6 empty] | 16
| 1 | 4 | < 5 | 4 | 2.2 | 7.2 | 5.9 | 2.3 | 3.1 | 10.6 | 0.7 | 1 | 1.7 | < 1.5 | 1.3094 | 1 | 1.4 | < 10 | 0.67 | 1.6 | 2 | 5 | 6 | 0.6 | < 10 | [6 empty] | 2.3602
| 1 | 5 | < 5 | 4.4 | 1.8 | 7.1 | 4.6 | 2.5 | 3.1 | 10.8 | 0.7 | 1 | 1.9 | < 1.5 | 1.5251 | 1 | 1.3 | < 10 | 0.66 | 1.6 | 4 | 5 | 6 | 0.6 | < 10 | [6 empty] | 13
| 1 | 6 | < 5 | 4.4 | 2.1 | 6.5 | 4.7 | 2.2 | 3.3 | 10.7 | 0.7 | 1 | 1.8 | < 1.5 | 1.6789 | 1 | 1.2 | < 10 | 0.64 | 1.4 | 3 | 5 | 5 | 0.6 | < 10
| Mean |  |  | 4.25 | 2.0167 | 6.9333 | 4.7833 | 2.5833 | 3.15 | 10.6667 | 0.7 | 1 | 1.8333 |  | 1.4765 | 1 | 1.3 |  | 0.66 | 1.5333 | 2.8333 | 5.1667 | 5.6667 | 0.61667
| Median |  |  | 4.3 | 2 | 7.05 | 4.65 | 2.4 | 3.1 | 10.65 | 0.7 | 1 | 1.85 |  | 1.4593 | 1 | 1.3 |  | 0.665 | 1.5 | 3 | 5 | 6 | 0.6
| Std Dev. |  |  | 0.16432 | 0.13292 | 0.27325 | 0.57764 | 0.55648 | 0.17607 | 0.33267 | 1.2162e-16 | 0 | 0.08165 |  | 0.17325 | 0 | 0.089443 |  | 0.016733 | 0.1633 | 0.75277 | 0.40825 | 0.5164 | 0.040825
| Rel.Std.Dev. |  |  | 3.866% | 6.591% | 3.941% | 12.08% | 21.54% | 5.589% | 3.119% | 1.737e-14% | 0% | 4.454% |  | 11.73% | 0% | 6.88% |  | 2.535% | 10.65% | 26.57% | 7.902% | 9.113% | 6.62%
| PDM3 |  |  | 80.07% | -14.55% | 193.8% | 102.7% | 9.455% | 33.46% | 351.9% | -70.34% | -57.63% | -22.32% |  | -37.44% | -57.63% | -44.92% |  | -72.04% | -35.03% | 20.05% | 118.9% | 140.1% | -73.87%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.19 | 1.71 | 0.88 | 0.03 | 0.25 | 3.15 | 0.69 | 0.58 | 0.26 | 0.67 | 0.39 | 0.58 | 0.46 | 0.96 | 0.71 | 0.37 | 0.13 | 1.03 | 1.22 | 0.73 | 0.96
| Analytical results for Sm in OREAS 627 (Indicative Value 3.01 ppm) | [30 empty] | Indicative
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.22 | [6 empty] | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*MS | [6 empty] | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.5g | [6 empty] | 2
| 1 | 1 | 1.01: 2.8; 1.02: 3.354; 1.03: 2.7938; 1.04: 3.3899 | [6 empty] | 1
| 1 | 2 | 1.01: 2.7; 1.02: 3.254; 1.03: 2.9328; 1.04: 3.3666 | [6 empty] | 8
| 1 | 3 | 1.01: 2.9; 1.02: 3.27; 1.03: 2.9586; 1.04: 3.369 | [6 empty] | 16
| 1 | 4 | 1.01: 2.8; 1.02: 3.18; 1.03: 2.9325; 1.04: 3.0704 | [6 empty] | 3.0104
| 1 | 5 | 1.01: 2.6; 1.02: 3.158; 1.03: 2.9403; 1.04: 3.0527 | [6 empty] | 14
| 1 | 6 | 1.01: 2.5; 1.02: 2.917; 1.03: 2.8299; 1.04: 3.1797
| Mean |  | 1.01: 2.7167; 1.02: 3.1888; 1.03: 2.898; 1.04: 3.2381
| Median |  | 1.01: 2.75; 1.02: 3.217; 1.03: 2.9327; 1.04: 3.2731
| Std Dev. |  | 1.01: 0.1472; 1.02: 0.15036; 1.03: 0.068339; 1.04: 0.15658
| Rel.Std.Dev. |  | 1.01: 5.418%; 1.02: 4.715%; 1.03: 2.358%; 1.04: 4.836%
| PDM3 |  | 1.01: -9.757%; 1.02: 5.928%; 1.03: -3.734%; 1.04: 7.563%
| Z-Score (Absolute) |  | 1.01: 1.3; 1.02: 0.58; 1.03: 0.58; 1.04: 0.77
| Analytical results for Sn in OREAS 627 (Certified Value 1.14 ppm) | [30 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [7 empty] | 2
| 1 | 1 | < 10 | 1.37 | 0.9 | 1.1 | 1.2 | 1 | 1 | 0.9 | 1.2 | < 2.5 | 1.0341 | < 5 | 1.4 | 1 | < 10 | 1.7 | 1.4 | 0.8 | 0.9 | 1 | < 10 | 1.2 | [7 empty] | 1
| 1 | 2 | < 10 | 1.23 | 0.9 | 1.1 | 1.2 | 1 | 1 | 0.9 | 1.14 | < 2.5 | 1.0682 | < 5 | 1.4 | 1.1 | < 10 | 1.6 | 1.4 | 0.9 | 0.9 | 1 | < 10 | 1.1 | [7 empty] | 12
| 1 | 3 | < 10 | 1.32 | 0.9 | 1.2 | 1.3 | 1 | 1 | 0.9 | 1.17 | < 2.5 | 1.2467 | < 5 | 1.4 | 1 | < 10 | 1.6 | 1.3 | 0.9 | 0.9 | 1 | < 10 | 1.1 | [7 empty] | 16
| 1 | 4 | < 10 | 1.26 | 0.9 | 1.1 | 1.1 | 1 | 1.1 | 0.9 | 1.21 | < 2.5 | 1.26 | < 5 | 1.4 | 1.1 | < 10 | 1.6 | 1.3 | 0.9 | 0.9 | 1.2 | < 10 | 1.1 | [7 empty] | 1.1392
| 1 | 5 | < 10 | 1.21 | 0.9 | 1.1 | 1.2 | 1 | 1.1 | 0.9 | 1.18 | < 2.5 | 1.1094 | < 5 | 1.4 | 1.1 | < 10 | 1.7 | 1.3 | 0.9 | 0.9 | 1.2 | < 10 | 1.1 | [7 empty] | 104
| 1 | 6 | < 10 | 1.48 | 1 | 1.1 | 1.3 | 1 | 1.1 | 0.9 | 1.08 | < 2.5 | 1.0943 | < 5 | 1.3 | 1.1 | < 10 | 1.6 | 1.3 | 0.8 | 1 | 1.1 | < 10 | 1.2
| Mean |  |  | 1.3117 | 0.91667 | 1.1167 | 1.2167 | 1 | 1.05 | 0.9 | 1.1633 |  | 1.1354 |  | 1.3833 | 1.0667 |  | 1.6333 | 1.3333 | 0.86667 | 0.91667 | 1.0833 |  | 1.1333
| Median |  |  | 1.29 | 0.9 | 1.1 | 1.2 | 1 | 1.05 | 0.9 | 1.175 |  | 1.1018 |  | 1.4 | 1.1 |  | 1.6 | 1.3 | 0.9 | 0.9 | 1.05 |  | 1.1
| Std Dev. |  |  | 0.10147 | 0.040825 | 0.040825 | 0.075277 | 0 | 0.054772 | 0 | 0.04761 |  | 0.09493 |  | 0.040825 | 0.05164 |  | 0.05164 | 0.05164 | 0.05164 | 0.040825 | 0.098319 |  | 0.05164
| Rel.Std.Dev. |  |  | 7.736% | 4.454% | 3.656% | 6.187% | 0% | 5.216% | 0% | 4.093% |  | 8.361% |  | 2.951% | 4.841% |  | 3.162% | 3.873% | 5.958% | 4.454% | 9.076% |  | 4.556%
| PDM3 |  |  | 15.14% | -19.53% | -1.977% | 6.801% | -12.22% | -7.83% | -21% | 2.119% |  | -0.3296% |  | 21.43% | -6.367% |  | 43.38% | 17.04% | -23.92% | -19.53% | -4.904% |  | -0.5145%
| Z-Score (Absolute) |  | 15.22 | 0.59 | 0.98 | 0.19 | 0.21 | NA | 0.45 | 1.04 | 0 | 0.34 | 0.11 | 5.3 | 0.87 | 0.38 | 15.22 | 1.86 | 0.67 | 1.18 | 0.98 | 0.32 | 15.22 | 0.12
Results from laboratory 1.07 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Sr in OREAS 627 (Certified Value 16.7 ppm) | [30 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 1
| 1 | 1 | 22 | 8.8 | 14.9 | 16.3 | 18.9 | 13.5 | 17.7 | 15.1 | 14.2 | 14.84 | 14.6 | 20 | 17.1213 | 13.83 | 16.68 | 31.3 | 18.4 | 14 | 21.45 | 19.6 | 16.3 | 21 | 16.5 | 17 | 14.8 | 19.4 | 31.4 |  |  | 1
| 1 | 2 | 20 | 8.3 | 15.5 | 15.7 | 17.3 | 14.3 | 17 | 14.7 | 14.1 | 14.82 | 15 | 19 | 17.1986 | 17.888 | 16.6675 | 30.8 | 18.8 | 13 | 21.51 | 19.4 | 15.9 | 20 | 16.6 | 17.8 | 14.7 | 19.3 | 28.7 |  |  | 21
| 1 | 3 | 22 | 8.2 | 15.2 | 15.5 | 19.6 | 14 | 16.5 | 15 | 14.6 | 14.37 | 14.3 | 20 | 17.8454 | 16.709 | 16.0305 | 31.2 | 18.9 | 14 | 21.47 | 19 | 16.1 | 20 | 16.8 | 17.6 | 14.4 | 19.2 | 29.7 |  |  | 16
| 1 | 4 | 22 | 9 | 15.6 | 15.2 | 16.1 | 13.2 | 17.2 | 16.2 | 14.2 | 13.94 | 14.4 | 20 | 17.4902 | 18.277 | 16.3887 | 30.4 | 18.6 | 14 | 21.53 | 19 | 15.9 | 20 | 16.1 | 17.3 | 14.7 | 20.2 | 30.6 |  |  | 16.6559
| 1 | 5 | 22 | 7.5 | 15.6 | 15.2 | 18.4 | 13.4 | 16.8 | 15.9 | 14.7 | 14.24 | 15.3 | 20 | 17.2776 | 17.418 | 16.9862 | 31.5 | 19.3 | 14 | 21.54 | 19.8 | 15.8 | 20 | 16.8 | 17.4 | 13.5 | 19 | 23.1 |  |  | 105
| 1 | 6 | 22 | 7.4 | 15.4 | 15.6 | 17.2 | 13.4 | 16.8 | 14.7 | 14.4 | 13.81 | 14 | 20 | 17.3303 | 17.391 | 15.8531 | 31.4 | 18.7 | 14 | 21.44 | 19.1 | 15.5 | 20 | 16.9 | 17.7 | 15 | 19.8 | 31.1
| Mean |  | 21.6667 | 8.2 | 15.3667 | 15.5833 | 17.9167 | 13.6333 | 17 | 15.2667 | 14.3667 | 14.3367 | 14.6 | 19.8333 | 17.3772 | 16.9188 | 16.4343 | 31.1 | 18.7833 | 13.8333 | 21.49 | 19.3167 | 15.9167 | 20.1667 | 16.6167 | 17.4667 | 14.5167 | 19.4833 | 29.1
| Median |  | 22 | 8.25 | 15.45 | 15.55 | 17.85 | 13.45 | 16.9 | 15.05 | 14.3 | 14.305 | 14.5 | 20 | 17.304 | 17.4045 | 16.5281 | 31.25 | 18.75 | 14 | 21.49 | 19.25 | 15.9 | 20 | 16.7 | 17.5 | 14.7 | 19.35 | 30.15
| Std Dev. |  | 0.8165 | 0.65422 | 0.27325 | 0.40702 | 1.2828 | 0.42269 | 0.41473 | 0.63456 | 0.24221 | 0.43177 | 0.47749 | 0.40825 | 0.26133 | 1.6026 | 0.42948 | 0.41952 | 0.30605 | 0.40825 | 0.042426 | 0.33714 | 0.27142 | 0.40825 | 0.29269 | 0.29439 | 0.53448 | 0.44008 | 3.1003
| Rel.Std.Dev. |  | 3.768% | 7.978% | 1.778% | 2.612% | 7.16% | 3.1% | 2.44% | 4.157% | 1.686% | 3.012% | 3.271% | 2.058% | 1.504% | 9.472% | 2.613% | 1.349% | 1.629% | 2.951% | 0.1974% | 1.745% | 1.705% | 2.024% | 1.761% | 1.685% | 3.682% | 2.259% | 10.65%
| PDM3 |  | 30.08% | -50.77% | -7.74% | -6.44% | 7.57% | -18.15% | 2.066% | -8.341% | -13.74% | -13.92% | -12.34% | 19.08% | 4.331% | 1.579% | -1.33% | 86.72% | 12.77% | -16.95% | 29.02% | 15.97% | -4.438% | 21.08% | -0.2355% | 4.868% | -12.84% | 16.98% | 74.71%
| Z-Score (Absolute) |  | NA | 2.81 | 0.42 | 0.35 | 0.43 | 1 | 0.13 | 0.45 | 0.75 | 0.76 | 0.67 | NA | 0.25 | 0.1 | 0.06 | 4.84 | 0.72 | NA | 1.63 | 0.9 | 0.23 | NA | 0 | 0.28 | 0.7 | 0.96 | 4.17
Results from laboratories 1.01, 1.12, 1.18 and 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ta in OREAS 627 (Certified Value 0.009 ppm) | [30 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.15 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [11 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [11 empty] | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [11 empty] | 3
| 1 | 1 | 6 | < 0.05 | < 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | 0.04 | [11 empty] | 1
| 1 | 2 | 5 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | 0.01 | [11 empty] | 13
| 1 | 3 | 6 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | 0.01 | [11 empty] | 16
| 1 | 4 | 6 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | 0.05 | 0.04 | 0.04 | [11 empty] | 0.0087778
| 1 | 5 | 6 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | 0.01 | 0.02 | [11 empty] | 106
| 1 | 6 | 6 | < 0.05 | 0.01 | 0.01 | 0.01 | 0.01 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.1 | < 0.3 | < 0.05 | < 0.05 | < 0.05 | < 0.01 | 0.06
| Mean |  | 1.01: 5.8333; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.06: 0.01; 1.07: 0.01; 1.16: 0.05; 1.17: 0.025; 1.18: 0.03
| Median |  | 1.01: 6; 1.03: 0.01; 1.04: 0.01; 1.05: 0.01; 1.06: 0.01; 1.07: 0.01; 1.16: 0.05; 1.17: 0.025; 1.18: 0.03
| Std Dev. |  | 1.01: 0.40825; 1.03: 0; 1.04: 0; 1.05: 0; 1.06: 0; 1.07: 0; 1.17: 0.021213; 1.18: 0.02
| Rel.Std.Dev. |  | 1.01: 6.999%; 1.03: 0%; 1.04: 0%; 1.05: 0%; 1.06: 0%; 1.07: 0%; 1.17: 84.85%; 1.18: 66.67%
| PDM3 |  | 1.01: 66356%; 1.03: 13.92%; 1.04: 13.92%; 1.05: 13.92%; 1.06: 13.92%; 1.07: 13.92%; 1.16: 469.6%; 1.17: 184.8%; 1.18: 241.8%
| Z-Score (Absolute) |  | 261.11 | 0 | 0.71 | 0.67 | 0.67 | 0.67 | 0.79 | 0.9 | 0.9 | 0 | 0 | 1.12 | 5.62 | 0 | 0 | 0.19 | 0.6 | 0.22
| Analytical results for Tb in OREAS 627 (Certified Value 0.34 ppm) | [30 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.21; 1.05: 1.22; 1.06: 1.23; 1.07: 1.25; 1.08: 1.26 |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.2g; 1.05: 0.5g; 1.06: 0.25g; 1.07: 0.25g; 1.08: 0.25g |  |  | 3
| 1 | 1 | 1.01: 0.3; 1.02: 0.357; 1.03: 0.32389; 1.04: 0.5; 1.05: 0.35664; 1.06: 0.35; 1.07: 0.48; 1.08: 0.32 |  |  | 1
| 1 | 2 | 1.01: 0.3; 1.02: 0.34; 1.03: 0.33713; 1.04: 0.5; 1.05: 0.3648; 1.06: 0.36; 1.07: 0.48; 1.08: 0.31 |  |  | 36
| 1 | 3 | 1.01: 0.3; 1.02: 0.365; 1.03: 0.32901; 1.04: 0.5; 1.05: 0.3899; 1.06: 0.34; 1.07: 0.49; 1.08: 0.32 |  |  | 16
| 1 | 4 | 1.01: 0.3; 1.02: 0.34; 1.03: 0.33707; 1.04: 0.5; 1.05: 0.33369; 1.06: 0.35; 1.07: 0.49; 1.08: 0.33 |  |  | 0.33996
| 1 | 5 | 1.01: 0.3; 1.02: 0.354; 1.03: 0.32829; 1.04: 0.5; 1.05: 0.33658; 1.06: 0.35; 1.07: 0.48; 1.08: 0.32 |  |  | 107
| 1 | 6 | 1.01: 0.3; 1.02: 0.321; 1.03: 0.33079; 1.04: 0.5; 1.05: 0.34386; 1.06: 0.33; 1.07: 0.48; 1.08: 0.33
| Mean |  | 1.01: 0.3; 1.02: 0.34617; 1.03: 0.33103; 1.04: 0.5; 1.05: 0.35424; 1.06: 0.34667; 1.07: 0.48333; 1.08: 0.32167
| Median |  | 1.01: 0.3; 1.02: 0.347; 1.03: 0.3299; 1.04: 0.5; 1.05: 0.35025; 1.06: 0.35; 1.07: 0.48; 1.08: 0.32
| Std Dev. |  | 1.01: 0; 1.02: 0.015766; 1.03: 0.0052213; 1.04: 0; 1.05: 0.021123; 1.06: 0.010328; 1.07: 0.005164; 1.08: 0.0075277
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 4.554%; 1.03: 1.577%; 1.04: 0%; 1.05: 5.963%; 1.06: 2.979%; 1.07: 1.068%; 1.08: 2.34%
| PDM3 |  | 1.01: -11.75%; 1.02: 1.827%; 1.03: -2.625%; 1.04: 47.08%; 1.05: 4.203%; 1.06: 1.974%; 1.07: 42.18%; 1.08: -5.38%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.01; 1.03: 0.89; 1.04: NA; 1.05: 0.45; 1.06: 0.01; 1.07: 7.95; 1.08: 1.44
Results from laboratories 1.02 and 1.21 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 627 (Certified Value 0.11 ppm) | [30 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 15g | 01g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 2
| 1 | 1 | 9 | 0.11 | 2.5 | 0.14 | 0.11 | 0.12 | 0.11 | 0.13 | 0.09 | 0.11 | < 2 | < 5 | 0.09806 | 0.11 | 0.1 | < 10 | 0.09 | 0.19 | 0.11 | 0.1 | 0.12 | 0.13 | 0.15 | [6 empty] | 1
| 1 | 2 | 8 | 0.09 | 3.1 | 0.13 | 0.11 | 0.21 | 0.12 | 0.1 | 0.09 | 0.1 | < 2 | < 5 | 0.095837 | 0.12 | < 0.1 | < 10 | 0.13 | 0.14 | 0.12 | 0.11 | 0.12 | 0.12 | 0.14 | [6 empty] | 14
| 1 | 3 | 9 | 0.12 | 2 | 0.13 | 0.12 | 0.15 | 0.09 | 0.11 | 0.09 | 0.1 | < 2 | < 5 | 0.14068 | 0.12 | < 0.1 | < 10 | 0.16 | 0.18 | 0.09 | 0.11 | 0.11 | 0.14 | 0.11 | [6 empty] | 16
| 1 | 4 | 9 | 0.12 | 2.3 | 0.13 | 0.12 | 0.11 | 0.11 | 0.11 | 0.17 | 0.11 | < 2 | < 5 | 0.14904 | 0.11 | < 0.1 | < 10 | 0.07 | 0.16 | 0.13 | 0.1 | 0.11 | 0.13 | 0.16 | [6 empty] | 0.11419
| 1 | 5 | 9 | 0.1 | 2.7 | 0.14 | 0.11 | 0.14 | 0.11 | 0.13 | 0.12 | 0.11 | < 2 | < 5 | 0.11915 | 0.13 | 0.1 | < 10 | 0.07 | 0.17 | 0.11 | 0.12 | 0.13 | 0.12 | 0.14 | [6 empty] | 108
| 1 | 6 | 8 | 0.11 | 2.6 | 0.13 | 0.1 | 0.13 | 0.1 | 0.14 | 0.1 | 0.12 | < 2 | < 5 | 0.1312 | 0.1 | < 0.1 | < 10 | 0.13 | 0.14 | 0.11 | 0.11 | 0.12 | 0.12 | 0.17
| Mean |  | 8.6667 | 0.10833 | 2.5333 | 0.13333 | 0.11167 | 0.14333 | 0.10667 | 0.12 | 0.11 | 0.10833 |  |  | 0.12233 | 0.115 | 0.1 |  | 0.10833 | 0.16333 | 0.11167 | 0.10833 | 0.11833 | 0.12667 | 0.145
| Median |  | 9 | 0.11 | 2.55 | 0.13 | 0.11 | 0.135 | 0.11 | 0.12 | 0.095 | 0.11 |  |  | 0.12517 | 0.115 | 0.1 |  | 0.11 | 0.165 | 0.11 | 0.11 | 0.12 | 0.125 | 0.145
| Std Dev. |  | 0.5164 | 0.01169 | 0.37238 | 0.005164 | 0.0075277 | 0.03559 | 0.010328 | 0.015492 | 0.031623 | 0.0075277 |  |  | 0.022044 | 0.010488 | 0 |  | 0.037103 | 0.020656 | 0.013292 | 0.0075277 | 0.0075277 | 0.008165 | 0.020736
| Rel.Std.Dev. |  | 5.958% | 10.79% | 14.7% | 3.873% | 6.741% | 24.83% | 9.682% | 12.91% | 28.75% | 6.949% |  |  | 18.02% | 9.12% | 0% |  | 34.25% | 12.65% | 11.9% | 6.949% | 6.361% | 6.446% | 14.3%
| PDM3 |  | 7490% | -5.126% | 2119% | 16.77% | -2.207% | 25.52% | -6.586% | 5.091% | -3.667% | -5.126% |  |  | 7.129% | 0.7119% | -12.42% |  | -5.126% | 43.04% | -2.207% | -5.126% | 3.631% | 10.93% | 26.98%
| Z-Score (Absolute) |  | NA | 0.67 | 150.27 | 0.88 | 0.47 | 1.5 | 0.78 | 0.05 | 0.57 | 0.67 | 54.83 | 148.19 | 0.2 | 0.26 | 3.27 | 303.8 | 0.67 | 2.75 | 0.47 | 0.67 | 0.05 | 0.47 | 1.61
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Th in OREAS 627 (Certified Value 6.54 ppm) | [30 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 01g | 01g | 01g | 0.25g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 3
| 1 | 1 | 4.6 | 6.2 | 6.8 | 7.7 | 6.6 | 6.6 | 6.1 | 6.3 | 6.67 | 6.4 | 6.2525 | 8.2429 | < 5 | 7.97 | 5.3 | 7.68 | 3.917 | 7.5 | 4 | 6.2 | 6.4 | 6.4 | < 10 | [6 empty] | 1
| 1 | 2 | 4.5 | 6.1 | 6.7 | 7.2 | 6.8 | 6.9 | 6.3 | 6.4 | 6.47 | 6.3 | 6.3485 | 8.3483 | < 5 | 8.06 | 5.4 | 7.46 | 3.416 | 7 | 4 | 6.3 | 6 | 6.5 | < 10 | [6 empty] | 38
| 1 | 3 | 4.5 | 7 | 6.5 | 7.9 | 6.7 | 6.5 | 6.2 | 6.5 | 6.25 | 6.1 | 6.6286 | 8.0913 | < 5 | 8.19 | 5.3 | 7.2 | 3.3649 | 7.2 | 4 | 6.5 | 6.4 | 6.3 | < 10 | [6 empty] | 16
| 1 | 4 | 4.7 | 6.9 | 6.4 | 6 | 6.8 | 6.6 | 6.3 | 6.4 | 6.15 | 6.4 | 6.6113 | 8.2442 | < 5 | 8.2 | 5.3 | 7.26 | 3.6134 | 7 | 4 | 6.4 | 6.4 | 6.4 | < 10 | [6 empty] | 6.5392
| 1 | 5 | 4.3 | 7 | 6.5 | 7.6 | 6.6 | 6.4 | 6.2 | 6.4 | 6.27 | 6.5 | 6.5363 | 8.2697 | < 5 | 8.11 | 5.4 | 7.56 | 3.5108 | 7.2 | 4 | 6.1 | 6 | 6.6 | < 10 | [6 empty] | 109
| 1 | 6 | 4.2 | 6.5 | 6.4 | 7.2 | 6.7 | 6.8 | 6.4 | 6.4 | 6.07 | 6.3 | 6.3928 | 8.0998 | < 5 | 8.21 | 5.3 | 7.45 | 3.6497 | 7 | 3 | 6.2 | 6.3 | 6.6 | < 10
| Mean |  | 4.4667 | 6.6167 | 6.55 | 7.2667 | 6.7 | 6.6333 | 6.25 | 6.4 | 6.3133 | 6.3333 | 6.4617 | 8.216 |  | 8.1233 | 5.3333 | 7.435 | 3.5786 | 7.15 | 3.8333 | 6.2833 | 6.25 | 6.4667
| Median |  | 4.5 | 6.7 | 6.5 | 7.4 | 6.7 | 6.6 | 6.25 | 6.4 | 6.26 | 6.35 | 6.4646 | 8.2436 |  | 8.15 | 5.3 | 7.455 | 3.5621 | 7.1 | 4 | 6.25 | 6.35 | 6.45
| Std Dev. |  | 0.18619 | 0.40702 | 0.16432 | 0.6802 | 0.089443 | 0.18619 | 0.10488 | 0.063246 | 0.2207 | 0.13663 | 0.15307 | 0.10096 |  | 0.095429 | 0.05164 | 0.18019 | 0.19874 | 0.19748 | 0.40825 | 0.1472 | 0.19748 | 0.12111
| Rel.Std.Dev. |  | 4.168% | 6.151% | 2.509% | 9.36% | 1.335% | 2.807% | 1.678% | 0.9882% | 3.496% | 2.157% | 2.369% | 1.229% |  | 1.175% | 0.9682% | 2.424% | 5.553% | 2.762% | 10.65% | 2.343% | 3.16% | 1.873%
| PDM3 |  | -31.69% | 1.185% | 0.1656% | 11.13% | 2.46% | 1.44% | -4.422% | -2.128% | -3.454% | -3.148% | -1.185% | 25.64% |  | 24.23% | -18.44% | 13.7% | -45.27% | 9.341% | -41.38% | -3.912% | -4.422% | -1.109%
| Z-Score (Absolute) |  | 5.62 | 0.53 | 0.34 | 2.39 | 0.77 | 0.58 | 0.52 | 0.09 | 0.34 | 0.28 | 0.09 | 5.1 | 11.24 | 4.84 | 3.14 | 2.87 | 8.16 | 2.06 | NA | 0.42 | 0.52 | 0.1 | 4.09
Results from laboratory 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 627 (Certified Value 0.029 wt.%) | [30 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 3
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  |  | 3
| 1 | 1 | 0.03 | 0.05 | 0.018 | 0.027 | 0.033 | 0.026 | 0.026 | 0.027 | 0.02 | 0.0322 | 0.024 | 0.0416 | 0.029134 | < 0.1 | 0.049342 | 0.04 | 0.0378 | 0.028 | 0.05 | 0.04 | 0.03 | 0.03 | 0.022 | 0.036 | 0.022 |  |  |  |  | 1
| 1 | 2 | 0.03 | 0.04 | 0.021 | 0.027 | 0.03 | 0.026 | 0.026 | 0.027 | 0.02 | 0.0305 | 0.024 | 0.0391 | 0.029949 | < 0.1 | 0.049221 | 0.04 | 0.0384 | 0.028 | 0.05 | 0.03 | 0.03 | 0.03 | 0.024 | 0.035 | 0.02 |  |  |  |  | 39
| 1 | 3 | 0.03 | 0.05 | 0.021 | 0.027 | 0.034 | 0.028 | 0.024 | 0.027 | 0.02 | 0.0296 | 0.022 | 0.0409 | 0.029254 | < 0.1 | 0.048509 | 0.04 | 0.0386 | 0.028 | 0.05 | 0.03 | 0.03 | 0.03 | 0.022 | 0.036 | 0.022 |  |  |  |  | 16
| 1 | 4 | 0.03 | 0.05 | 0.022 | 0.026 | 0.031 | 0.027 | 0.026 | 0.026 | 0.02 | 0.0289 | 0.021 | 0.0407 | 0.03009 | < 0.1 | 0.048763 | 0.04 | 0.0392 | 0.028 | 0.05 | 0.03 | 0.03 | 0.02 | 0.023 | 0.036 | 0.022 |  |  |  |  | 0.028857
| 1 | 5 | 0.03 | 0.04 | 0.023 | 0.027 | 0.032 | 0.027 | 0.026 | 0.027 | 0.02 | 0.0296 | 0.024 | 0.0412 | 0.029709 | < 0.1 | 0.048464 | 0.05 | 0.0386 | 0.028 | 0.05 | 0.03 | 0.03 | 0.02 | 0.02 | 0.036 | 0.019 |  |  |  |  | 110
| 1 | 6 | 0.03 | 0.05 | 0.021 | 0.027 | 0.032 | 0.028 | 0.025 | 0.027 | 0.02 | 0.0261 | 0.021 | 0.0418 | 0.029498 | < 0.1 | 0.049519 | 0.04 | 0.0393 | 0.028 | 0.05 | 0.03 | 0.03 | 0.03 | 0.029 | 0.037 | 0.021
| Mean |  | 0.03 | 0.046667 | 0.021 | 0.026833 | 0.032 | 0.027 | 0.0255 | 0.026833 | 0.02 | 0.029483 | 0.022667 | 0.040883 | 0.029606 |  | 0.04897 | 0.041667 | 0.03865 | 0.028 | 0.05 | 0.031667 | 0.03 | 0.026667 | 0.023333 | 0.036 | 0.021
| Median |  | 0.03 | 0.05 | 0.021 | 0.027 | 0.032 | 0.027 | 0.026 | 0.027 | 0.02 | 0.0296 | 0.023 | 0.04105 | 0.029604 |  | 0.048992 | 0.04 | 0.0386 | 0.028 | 0.05 | 0.03 | 0.03 | 0.03 | 0.0225 | 0.036 | 0.0215
| Std Dev. |  | 0 | 0.005164 | 0.0016733 | 0.00040825 | 0.0014142 | 0.00089443 | 0.00083666 | 0.00040825 | 0 | 0.0020114 | 0.0015055 | 0.00096626 | 0.00037967 |  | 0.0004504 | 0.0040825 | 0.00055045 | 3.8006e-18 | 7.6012e-18 | 0.0040825 | 0 | 0.005164 | 0.0030768 | 0.00063246 | 0.0012649
| Rel.Std.Dev. |  | 0% | 11.07% | 7.968% | 1.521% | 4.419% | 3.313% | 3.281% | 1.521% | 0% | 6.822% | 6.642% | 2.363% | 1.282% |  | 0.9197% | 9.798% | 1.424% | 1.357e-14% | 1.52e-14% | 12.89% | 0% | 19.36% | 13.19% | 1.757% | 6.023%
| PDM3 |  | 3.962% | 61.72% | -27.23% | -7.012% | 10.89% | -6.434% | -11.63% | -7.012% | -30.69% | 2.172% | -21.45% | 41.68% | 2.596% |  | 69.7% | 44.39% | 33.94% | -2.969% | 73.27% | 9.738% | 3.962% | -7.589% | -19.14% | 24.75% | -27.23%
| Z-Score (Absolute) |  | 0.04 | 1.83 | 0.93 | 0.3 | 0.26 | 0.28 | 0.44 | 0.3 | 1.03 | 0.01 | 0.75 | 1.21 | 0 | 2.19 | 2.08 | 1.3 | 0.97 | 0.17 | 2.19 | 0.22 | 0.04 | 0.32 | 0.67 | 0.69 | 0.93
| Analytical results for Tl in OREAS 627 (Certified Value 5.28 ppm) | [30 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.5g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g |  |  |  | 3
| 1 | 1 | < 10 | 5.45 | 5.1 | 4.86 | 6.02 | 5.21 | 5.24 | 4.99 | 5.25 | 6.05 | 5.53 | < 5 | 4.5361 | 17.42 | 5.2121 | < 5 | 6.4 | 3.18 | 5.066 | 6.57 | 5.55 | < 5 | 4.91 | 5.28 | 5.44 | < 5 |  |  |  | 1
| 1 | 2 | < 10 | 5.29 | 5.1 | 4.9 | 5.6 | 5.56 | 5.23 | 5.12 | 5.31 | 5.79 | 5.43 | < 5 | 4.6554 | 20.26 | 5.1462 | < 5 | 6.39 | 3.11 | 5.065 | 6.38 | 5.09 | < 5 | 5.02 | 5.11 | 5.29 | < 5 |  |  |  | 0
| 1 | 3 | < 10 | 5.36 | 5.4 | 4.66 | 6.26 | 5.27 | 5.11 | 5.08 | 5.35 | 5.78 | 5.24 | < 5 | 4.6326 | 16.42 | 5.1336 | < 5 | 6.41 | 3.15 | 5.046 | 6.16 | 5.42 | 5 | 5.09 | 5.36 | 5.38 | < 5 |  |  |  | 16
| 1 | 4 | < 10 | 5.36 | 5.4 | 4.81 | 5.14 | 5.36 | 5.08 | 5.27 | 5.37 | 5.66 | 5.39 | < 5 | 4.6603 | 18.56 | 5.1767 | < 5 | 6.4 | 3.18 | 5.037 | 6.4 | 5.3 | < 5 | 4.97 | 5.37 | 5.29 | < 5 |  |  |  | 5.2815
| 1 | 5 | < 10 | 5.16 | 5.5 | 4.88 | 5.94 | 5.33 | 5.08 | 5.17 | 5.36 | 5.82 | 5.6 | < 5 | 4.6107 | 15.48 | 5.1996 | < 5 | 6.45 | 3.23 | 5.193 | 6.6 | 5.32 | 5 | 5.02 | 5.16 | 5.17 | < 5 |  |  |  | 111
| 1 | 6 | < 10 | 5.34 | 5.3 | 4.86 | 5.86 | 5.44 | 5.1 | 5.31 | 5.27 | 5.46 | 5.19 | < 5 | 4.5315 | 15.51 | 5.1878 | < 5 | 6.41 | 3.17 | 4.968 | 6.38 | 5.37 | 5 | 5.14 | 5.25 | 5.29 | < 5
| Mean |  |  | 5.3267 | 5.3 | 4.8283 | 5.8033 | 5.3617 | 5.14 | 5.1567 | 5.3183 | 5.76 | 5.3967 |  | 4.6044 | 17.275 | 5.176 |  | 6.41 | 3.17 | 5.0625 | 6.415 | 5.3417 | 5 | 5.025 | 5.255 | 5.31
| Median |  |  | 5.35 | 5.35 | 4.86 | 5.9 | 5.345 | 5.105 | 5.145 | 5.33 | 5.785 | 5.41 |  | 4.6216 | 16.92 | 5.1823 |  | 6.405 | 3.175 | 5.0555 | 6.39 | 5.345 | 5 | 5.02 | 5.265 | 5.29
| Std Dev. |  |  | 0.096678 | 0.16733 | 0.087731 | 0.38955 | 0.12481 | 0.074565 | 0.11961 | 0.049967 | 0.19442 | 0.15971 |  | 0.057521 | 1.8789 | 0.030611 |  | 0.020976 | 0.039497 | 0.073351 | 0.15871 | 0.15224 | 0 | 0.082158 | 0.10483 | 0.092304
| Rel.Std.Dev. |  |  | 1.815% | 3.157% | 1.817% | 6.712% | 2.328% | 1.451% | 2.32% | 0.9395% | 3.375% | 2.959% |  | 1.249% | 10.88% | 0.5914% |  | 0.3272% | 1.246% | 1.449% | 2.474% | 2.85% | 0% | 1.635% | 1.995% | 1.738%
| PDM3 |  |  | 0.855% | 0.3501% | -8.58% | 9.88% | 1.518% | -2.679% | -2.364% | 0.6972% | 9.06% | 2.18% |  | -12.82% | 227.1% | -1.998% |  | 21.37% | -39.98% | -4.147% | 21.46% | 1.139% | -5.33% | -4.857% | -0.5019% | 0.5394%
| Z-Score (Absolute) |  | 0.75 | 0.21 | 0.13 | 1.25 | 1.61 | 0.31 | 0.34 | 0.29 | 0.19 | 1.48 | 0.42 | 8.08 | 1.91 | 35.24 | 0.23 | 8.08 | 3.39 | 6.11 | 0.56 | 3.4 | 0.25 | NA | 0.67 | 0 | 0.16 | 8.08
Results from laboratory 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Tm in OREAS 627 (Indicative Value 0.073 ppm) | [30 empty] | Indicative
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13 |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS |  |  |  | 3
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g |  |  |  | 3
| 1 | 1 | 1.01: < 0.1; 1.02: 0.084; 1.03: 0.067628 |  |  |  | 1
| 1 | 2 | 1.01: < 0.1; 1.02: 0.077; 1.03: 0.066435 |  |  |  | 9
| 1 | 3 | 1.01: < 0.1; 1.02: 0.079; 1.03: 0.066596 |  |  |  | 16
| 1 | 4 | 1.01: < 0.1; 1.02: 0.082; 1.03: 0.066366 |  |  |  | 0.07256
| 1 | 5 | 1.01: < 0.1; 1.02: 0.077; 1.03: 0.063004 |  |  |  | 15
| 1 | 6 | 1.01: < 0.1; 1.02: 0.079; 1.03: 0.062688
| Mean |  | 1.02: 0.079667; 1.03: 0.065453
| Median |  | 1.02: 0.079; 1.03: 0.066401
| Std Dev. |  | 1.02: 0.0028048; 1.03: 0.0020726
| Rel.Std.Dev. |  | 1.02: 3.521%; 1.03: 3.167%
| PDM3 |  | 1.02: 9.795%; 1.03: -9.795%
| Z-Score (Absolute) |  | 1.01: 0.73; 1.02: 0.67; 1.03: 0
| Analytical results for U in OREAS 627 (Certified Value 2.52 ppm) | [30 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.17 | 1.18 | 1.20 | 1.21 | 1.22 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.30 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | [6 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 0.2g | 01g | 01g | 0.25g | 0.5g | 0.2g | 0.5g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | [6 empty] | 3
| 1 | 1 | 2.6 | 2.2 | 2.16 | 2.62 | 2.39 | 2.48 | 2.36 | 2.18 | 3.02 | 2.5 | 2.1735 | 2.7738 | 3.09 | 2.3 | 2.62 | 2.96 | 1.4311 | 3.05 | < 8 | 2.34 | 2.29 | 2.49 | 2.37 | [6 empty] | 1
| 1 | 2 | 2.6 | 2.3 | 2.29 | 2.43 | 2.51 | 2.52 | 2.43 | 2.24 | 2.89 | 2.4 | 2.407 | 2.7912 | 2.98 | 2.3 | 2.61 | 2.92 | 1.4955 | 2.69 | < 8 | 2.4 | 2.19 | 2.51 | 2.31 | [6 empty] | 41
| 1 | 3 | 2.6 | 2.5 | 2.18 | 2.73 | 2.55 | 2.4 | 2.38 | 2.31 | 2.89 | 2.3 | 2.2664 | 2.829 | 2.97 | 2.3 | 2.67 | 2.85 | 1.5842 | 2.81 | < 8 | 2.52 | 2.34 | 2.42 | 2.42 | [6 empty] | 16
| 1 | 4 | 2.7 | 2.4 | 2.2 | 2.5 | 2.48 | 2.45 | 2.48 | 2.2 | 2.87 | 2.4 | 2.252 | 2.9125 | 3.01 | 2.3 | 2.66 | 2.88 | 1.3116 | 2.73 | < 8 | 2.49 | 2.3 | 2.47 | 2.35 | [6 empty] | 2.5185
| 1 | 5 | 2.5 | 2.4 | 2.26 | 2.53 | 2.44 | 2.45 | 2.4 | 2.21 | 2.9 | 2.5 | 2.2012 | 2.9255 | 3.04 | 2.4 | 2.68 | 2.98 | 1.2967 | 2.81 | < 8 | 2.46 | 2.22 | 2.36 | 2.42 | [6 empty] | 112
| 1 | 6 | 2.6 | 2.3 | 2.23 | 2.47 | 2.47 | 2.66 | 2.43 | 2.2 | 2.95 | 2.3 | 2.2052 | 2.9145 | 3.01 | 2.3 | 2.44 | 2.89 | 1.358 | 2.84 | < 8 | 2.44 | 2.27 | 2.55 | 2.27
| Mean |  | 2.6 | 2.35 | 2.22 | 2.5467 | 2.4733 | 2.4933 | 2.4133 | 2.2233 | 2.92 | 2.4 | 2.2509 | 2.8577 | 3.0167 | 2.3167 | 2.6133 | 2.9133 | 1.4128 | 2.8217 |  | 2.4417 | 2.2683 | 2.4667 | 2.3567
| Median |  | 2.6 | 2.35 | 2.215 | 2.515 | 2.475 | 2.465 | 2.415 | 2.205 | 2.895 | 2.4 | 2.2286 | 2.8707 | 3.01 | 2.3 | 2.64 | 2.905 | 1.3945 | 2.81 |  | 2.45 | 2.28 | 2.48 | 2.36
| Std Dev. |  | 0.063246 | 0.10488 | 0.049396 | 0.11039 | 0.055377 | 0.090701 | 0.04274 | 0.046762 | 0.055857 | 0.089443 | 0.083862 | 0.068015 | 0.043665 | 0.040825 | 0.089368 | 0.049666 | 0.11249 | 0.12529 |  | 0.064627 | 0.054924 | 0.067725 | 0.059889
| Rel.Std.Dev. |  | 2.433% | 4.463% | 2.225% | 4.335% | 2.239% | 3.638% | 1.771% | 2.103% | 1.913% | 3.727% | 3.726% | 2.38% | 1.447% | 1.762% | 3.42% | 1.705% | 7.962% | 4.44% |  | 2.647% | 2.421% | 2.746% | 2.541%
| PDM3 |  | 3.237% | -6.69% | -11.85% | 1.119% | -1.793% | -0.9985% | -4.175% | -11.72% | 15.94% | -4.704% | -10.63% | 13.47% | 19.78% | -8.013% | 3.766% | 15.68% | -43.9% | 12.04% |  | -3.05% | -9.932% | -2.057% | -6.425%
| Z-Score (Absolute) |  | 0.6 | 0.52 | 1.11 | 0.36 | 0.03 | 0.12 | 0.24 | 1.09 | 2.04 | 0.3 | 0.97 | 1.76 | 2.47 | 0.67 | 0.66 | 2.01 | 4.74 | 1.6 | 6.89 | 0.11 | 0.89 | 0 | 0.49
| Analytical results for V in OREAS 627 (Certified Value 8.42 ppm) | [30 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 0.5g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  | 3
| 1 | 1 | 10 | 9 | 8 | 8 | 9 | 8 | 8 | 8 | 6.9 | 9 | 9 | 10 | 7.7724 | 8.62 | 8.3647 | 18.7 | 12 | 7 | < 10 | 12 | 9 | 9 | 9 | 13 | 8 | 5 | 8 |  |  | 1
| 1 | 2 | 9 | 9 | 9 | 8 | 8 | 8 | 8 | 8 | 6.8 | 9 | 8 | 10 | 8.1527 | 9.05 | 8.4159 | 17.3 | 12 | 7 | < 10 | 12 | 9 | 8 | 8 | 13 | 8 | 6 | 8 |  |  | 23
| 1 | 3 | 10 | 9 | 9 | 8 | 9 | 8 | 8 | 8 | 7.1 | 9 | 8 | 10 | 8.1797 | 7.53 | 8.2549 | 18.1 | 13 | 8 | < 10 | 12 | 9 | 8 | 8 | 14 | 8 | 6 | 8 |  |  | 16
| 1 | 4 | 10 | 9 | 9 | 8 | 9 | 9 | 8 | 8 | 6.9 | 8 | 9 | 10 | 8.0894 | 8.55 | 7.9658 | 17 | 13 | 7 | < 10 | 12 | 9 | 9 | 9 | 13 | 9 | 6 | 7 |  |  | 8.4201
| 1 | 5 | 10 | 8 | 10 | 8 | 8 | 8 | 8 | 8 | 7.1 | 8 | 9 | 10 | 7.9754 | 8.36 | 8.01 | 19.3 | 12 | 8 | < 10 | 12 | 9 | 9 | 9 | 13 | 8 | 6 | 7 |  |  | 113
| 1 | 6 | 10 | 9 | 9 | 8 | 8 | 9 | 8 | 8 | 7.3 | 7 | 9 | 10 | 7.7324 | 8.82 | 7.8434 | 17.9 | 12 | 7 | < 10 | 12 | 9 | 9 | 9 | 13 | 9 | 5 | 7
| Mean |  | 9.8333 | 8.8333 | 9 | 8 | 8.5 | 8.3333 | 8 | 8 | 7.0167 | 8.3333 | 8.6667 | 10 | 7.9837 | 8.4883 | 8.1425 | 18.05 | 12.3333 | 7.3333 |  | 12 | 9 | 8.6667 | 8.6667 | 13.1667 | 8.3333 | 5.6667 | 7.5
| Median |  | 10 | 9 | 9 | 8 | 8.5 | 8 | 8 | 8 | 7 | 8.5 | 9 | 10 | 8.0324 | 8.585 | 8.1325 | 18 | 12 | 7 |  | 12 | 9 | 9 | 9 | 13 | 8 | 6 | 7.5
| Std Dev. |  | 0.40825 | 0.40825 | 0.63246 | 0 | 0.54772 | 0.5164 | 0 | 0 | 0.18348 | 0.8165 | 0.5164 | 0 | 0.19287 | 0.52564 | 0.23453 | 0.85732 | 0.5164 | 0.5164 |  | 0 | 0 | 0.5164 | 0.5164 | 0.40825 | 0.5164 | 0.5164 | 0.54772
| Rel.Std.Dev. |  | 4.152% | 4.622% | 7.027% | 0% | 6.444% | 6.197% | 0% | 0% | 2.615% | 9.798% | 5.958% | 0% | 2.416% | 6.192% | 2.88% | 4.75% | 4.187% | 7.042% |  | 0% | 0% | 5.958% | 5.958% | 3.101% | 6.197% | 9.113% | 7.303%
| PDM3 |  | 16.78% | 4.907% | 6.887% | -4.99% | 0.9485% | -1.031% | -4.99% | -4.99% | -16.67% | -1.031% | 2.928% | 18.76% | -5.183% | 0.81% | -3.298% | 114.4% | 46.47% | -12.91% |  | 42.52% | 6.887% | 2.928% | 2.928% | 56.37% | -1.031% | -32.7% | -10.93%
| Z-Score (Absolute) |  | 1.8 | 0.46 | 0.68 | 0.65 | 0.02 | 0.21 | 0.65 | 0.65 | 1.97 | 0.21 | 0.24 | 2.02 | 0.67 | 0 | 0.46 | 12.78 | 5.14 | 1.54 | 4.66 | 4.69 | 0.68 | 0.24 | 0.24 | 6.25 | 0.21 | 3.77 | 1.32
| Analytical results for W in OREAS 627 (Certified Value 4.29 ppm) | [30 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS |  |  |  |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 15g | 01g | 0.2g | 01g | 01g | 01g | 0.25g | 0.2g | 0.25g | 0.5g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g |  |  |  |  | 3
| 1 | 1 | < 5 | 4.6 | 3.5 | 3.7 | 4.61 | 4.05 | 4.12 | 3.97 | 4.3 | 4.34 | 3.3 | 179 | < 10 | 4.1357 | < 5 | 4.5 | 4.3 | 4.5 | 5 | 4 | 4.5 | 4.4 | 4.01 | < 10 | 5.23 |  |  |  |  | 1
| 1 | 2 | < 5 | 4.6 | 3.5 | 3.75 | 4.31 | 4.21 | 3.98 | 4.04 | 4.37 | 4.29 | 3.3 | > 200 | < 10 | 4.0879 | < 5 | 4.4 | 4.3 | 4.6 | 4.6 | 4 | 4.6 | 4.4 | 4.09 | < 10 | 4.25 |  |  |  |  | 43
| 1 | 3 | < 5 | 5.1 | 3.7 | 3.7 | 4.63 | 4.18 | 3.93 | 3.84 | 4.36 | 4.15 | 3.1 | 197 | < 10 | 4.0972 | < 5 | 4.5 | 4.3 | 4.3 | 4.9 | 4 | 4.6 | 4.6 | 4.05 | < 10 | 4.89 |  |  |  |  | 16
| 1 | 4 | < 5 | 4.6 | 3.6 | 3.67 | 4.13 | 4.08 | 4.01 | 4.02 | 4.36 | 4.28 | 3.1 | 169 | < 10 | 4.0529 | < 5 | 4.5 | 4.3 | 4.4 | 4.6 | 4 | 4.7 | 4.5 | 4.21 | < 10 | 5.24 |  |  |  |  | 4.2946
| 1 | 5 | < 5 | 4.3 | 3.7 | 3.76 | 4.58 | 3.93 | 4.18 | 3.99 | 4.41 | 4.27 | 3.2 | 189 | < 10 | 4.1767 | < 5 | 4.4 | 4.4 | 4.5 | 4.8 | 4 | 4.6 | 4.3 | 4.15 | < 10 | 4.7 |  |  |  |  | 114
| 1 | 6 | < 5 | 4.4 | 3.6 | 3.74 | 4.39 | 4.05 | 4.15 | 4.01 | 4.32 | 4.11 | 2.9 | 177 | < 10 | 4.1476 | < 5 | 4.4 | 4.5 | 4.5 | 4.6 | 3 | 4.6 | 4.5 | 4.1 | < 10 | 5.16
| Mean |  |  | 4.6 | 3.6 | 3.72 | 4.4417 | 4.0833 | 4.0617 | 3.9783 | 4.3533 | 4.24 | 3.15 | 182.2 |  | 4.1163 |  | 4.45 | 4.35 | 4.4667 | 4.75 | 3.8333 | 4.6 | 4.45 | 4.1017 |  | 4.9117
| Median |  |  | 4.6 | 3.6 | 3.72 | 4.485 | 4.065 | 4.065 | 4 | 4.36 | 4.275 | 3.15 | 179 |  | 4.1164 |  | 4.45 | 4.3 | 4.5 | 4.7 | 4 | 4.6 | 4.45 | 4.095 |  | 5.025
| Std Dev. |  |  | 0.27568 | 0.089443 | 0.035214 | 0.20004 | 0.10113 | 0.10187 | 0.071949 | 0.038816 | 0.089443 | 0.15166 | 10.9179 |  | 0.045152 |  | 0.054772 | 0.083666 | 0.10328 | 0.17607 | 0.40825 | 0.063246 | 0.10488 | 0.07111 |  | 0.3883
| Rel.Std.Dev. |  |  | 5.993% | 2.485% | 0.9466% | 4.504% | 2.477% | 2.508% | 1.809% | 0.8916% | 2.109% | 4.815% | 5.992% |  | 1.097% |  | 1.231% | 1.923% | 2.312% | 3.707% | 10.65% | 1.375% | 2.357% | 1.734% |  | 7.906%
| PDM3 |  |  | 7.112% | -16.17% | -13.38% | 3.425% | -4.919% | -5.423% | -7.364% | 1.368% | -1.271% | -26.65% | 4143% |  | -4.15% |  | 3.619% | 1.291% | 4.007% | 10.6% | -10.74% | 7.112% | 3.619% | -4.492% |  | 14.37%
| Z-Score (Absolute) |  | 4.47 | 0.6 | 1.82 | 1.53 | 0.22 | 0.65 | 0.7 | 0.9 | 0 | 0.27 | 2.9 | 429.58 | 1.57 | 0.57 | 4.47 | 0.24 | 0 | 0.28 | 0.96 | NA | 0.6 | 0.24 | 0.6 | 1.57 | 1.35
Results from laboratory 1.24 were removed due to their 1 ppm reading resolution.
| Analytical results for Y in OREAS 627 (Certified Value 6.27 ppm) | [30 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [8 empty] | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [8 empty] | 3
| 1 | 1 | 7 | 6.68 | 6.12 | 7.25 | 5.94 | 6.04 | 5.81 | 5.03 | 6.86 | 6.1347 | 6.8702 | 7.56 | 5.4 | < 10 | 8.3 | 6.82 | 5.55 | 5.97 | 6.34 | 6.5 | 6.31 | [8 empty] | 1
| 1 | 2 | 6 | 6.6 | 6.35 | 6.76 | 6.1 | 6.09 | 5.98 | 5.05 | 6.66 | 6.337 | 6.805 | 7.5 | 5.6 | < 10 | 8 | 6.63 | 5.79 | 6.13 | 6.38 | 6.4 | 5.93 | [8 empty] | 44
| 1 | 3 | 7 | 6.85 | 6.33 | 7.41 | 6.04 | 5.78 | 5.77 | 5.18 | 6.8 | 6.2914 | 6.5585 | 7.44 | 5.6 | < 10 | 8 | 6.62 | 5.91 | 6 | 6.47 | 6.4 | 6.14 | [8 empty] | 16
| 1 | 4 | 7 | 6.59 | 6.32 | 6.44 | 5.75 | 5.93 | 5.8 | 4.98 | 6.71 | 6.4046 | 6.6413 | 7.56 | 5.5 | < 10 | 7.9 | 6.72 | 5.49 | 6.23 | 6.18 | 6.3 | 6.1 | [8 empty] | 6.2712
| 1 | 5 | 7 | 6.42 | 6.53 | 6.99 | 6.01 | 5.85 | 5.74 | 5.24 | 6.68 | 6.3803 | 6.6358 | 7.38 | 5.6 | < 10 | 8.1 | 6.82 | 5.76 | 6.14 | 6.35 | 6.3 | 6 | [8 empty] | 115
| 1 | 6 | 7 | 6.82 | 6.35 | 6.84 | 6.02 | 5.96 | 5.68 | 5.17 | 6.5 | 6.2593 | 6.7766 | 7.54 | 5.5 | < 10 | 8.1 | 6.5 | 5.54 | 6.18 | 5.96 | 6.6 | 6.1
| Mean |  | 6.8333 | 6.66 | 6.3333 | 6.9483 | 5.9767 | 5.9417 | 5.7967 | 5.1083 | 6.7017 | 6.3012 | 6.7146 | 7.4967 | 5.5333 |  | 8.0667 | 6.685 | 5.6733 | 6.1083 | 6.28 | 6.4167 | 6.0967
| Median |  | 7 | 6.64 | 6.34 | 6.915 | 6.015 | 5.945 | 5.785 | 5.11 | 6.695 | 6.3142 | 6.709 | 7.52 | 5.55 |  | 8.05 | 6.675 | 5.655 | 6.135 | 6.345 | 6.4 | 6.1
| Std Dev. |  | 0.40825 | 0.16012 | 0.13033 | 0.34971 | 0.12242 | 0.11548 | 0.10132 | 0.10226 | 0.12465 | 0.097784 | 0.12015 | 0.073121 | 0.08165 |  | 0.13663 | 0.12582 | 0.16955 | 0.10226 | 0.18276 | 0.1169 | 0.13003
| Rel.Std.Dev. |  | 5.974% | 2.404% | 2.058% | 5.033% | 2.048% | 1.944% | 1.748% | 2.002% | 1.86% | 1.552% | 1.789% | 0.9754% | 1.476% |  | 1.694% | 1.882% | 2.989% | 1.674% | 2.91% | 1.822% | 2.133%
| PDM3 |  | 8.964% | 6.2% | 0.9906% | 10.8% | -4.697% | -5.255% | -7.567% | -18.54% | 6.864% | 0.4783% | 7.07% | 19.54% | -11.77% |  | 28.63% | 6.598% | -9.534% | -2.597% | 0.1402% | 2.319% | -2.783%
| Z-Score (Absolute) |  | NA | 0.62 | 0.07 | 1.1 | 0.53 | 0.58 | 0.83 | 1.98 | 0.69 | 0.02 | 0.71 | 2.02 | 1.27 | 2.16 | 2.97 | 0.66 | 1.03 | 0.31 | 0.02 | 0.21 | 0.32
Results from laboratory 1.01 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 627 (Certified Value 0.44 ppm) | [30 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.21; 1.05: 1.22; 1.06: 1.23; 1.07: 1.25; 1.08: 1.26 |  |  |  |  | ppm
|  |  | 1.01: AR*OES/MS; 1.02: AR*MS; 1.03: AR*MS; 1.04: AR*OES/MS; 1.05: AR*MS; 1.06: AR*MS; 1.07: AR*OES/MS; 1.08: AR*OES/MS |  |  |  |  | 2
|  |  | 1.01: 0.5g; 1.02: 25g; 1.03: 0.2g; 1.04: 0.2g; 1.05: 0.5g; 1.06: 0.25g; 1.07: 0.25g; 1.08: 0.25g |  |  |  |  | 2
| 1 | 1 | 1.01: 0.5; 1.02: 0.509; 1.03: 0.39207; 1.04: 0.6; 1.05: 0.43045; 1.06: 0.5; 1.07: 0.4; 1.08: 0.4 |  |  |  |  | 1
| 1 | 2 | 1.01: 0.5; 1.02: 0.471; 1.03: 0.4049; 1.04: 0.6; 1.05: 0.42206; 1.06: 0.4; 1.07: 0.4; 1.08: 0.4 |  |  |  |  | 45
| 1 | 3 | 1.01: 0.5; 1.02: 0.499; 1.03: 0.3869; 1.04: 0.6; 1.05: 0.47291; 1.06: 0.5; 1.07: 0.5; 1.08: 0.4 |  |  |  |  | 16
| 1 | 4 | 1.01: 0.5; 1.02: 0.503; 1.03: 0.37972; 1.04: 0.6; 1.05: 0.39221; 1.06: 0.5; 1.07: 0.4; 1.08: 0.4 |  |  |  |  | 0.43899
| 1 | 5 | 1.01: 0.5; 1.02: 0.473; 1.03: 0.39017; 1.04: 0.6; 1.05: 0.39963; 1.06: 0.5; 1.07: 0.4; 1.08: 0.4 |  |  |  |  | 116
| 1 | 6 | 1.01: 0.5; 1.02: 0.433; 1.03: 0.37277; 1.04: 0.6; 1.05: 0.40575; 1.06: 0.4; 1.07: 0.4; 1.08: 0.4
| Mean |  | 1.01: 0.5; 1.02: 0.48133; 1.03: 0.38776; 1.04: 0.6; 1.05: 0.4205; 1.06: 0.46667; 1.07: 0.41667; 1.08: 0.4
| Median |  | 1.01: 0.5; 1.02: 0.486; 1.03: 0.38854; 1.04: 0.6; 1.05: 0.41391; 1.06: 0.5; 1.07: 0.4; 1.08: 0.4
| Std Dev. |  | 1.01: 0; 1.02: 0.028493; 1.03: 0.011029; 1.04: 0; 1.05: 0.029332; 1.06: 0.05164; 1.07: 0.040825; 1.08: 6.0809e-17
| Rel.Std.Dev. |  | 1.01: 0%; 1.02: 5.92%; 1.03: 2.844%; 1.04: 0%; 1.05: 6.976%; 1.06: 11.07%; 1.07: 9.798%; 1.08: 1.52e-14%
| PDM3 |  | 1.01: 13.9%; 1.02: 9.646%; 1.03: -11.67%; 1.04: 36.68%; 1.05: -4.211%; 1.06: 6.305%; 1.07: -5.085%; 1.08: -8.882%
| Z-Score (Absolute) |  | 1.01: 0.94; 1.02: 0.63; 1.03: 0.93; 1.04: 2.59; 1.05: 0.38; 1.06: 0.38; 1.07: 0.45; 1.08: 0.72
| Analytical results for Zn in OREAS 627 (Certified Value 10.07 wt.%) | [30 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 |  | wt.%
|  |  | AR*AAS | AR*OES | AR*AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/AAS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*AAS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*AAS | AR*OES | AR*OES | AR*OES/MS | AR*AAS | AR*OES | AR*OES | AR*MS |  | 2
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 01g | 01g | 0.2g | 0.5g | 0.1g | 01g | 01g | 0.25g | 0.25g | 0.5g | 0.25g | 0.25g | 01g | 0.25g | 0.25g | 0.4g | 0.25g | 0.2g |  | 3
| 1 | 1 | 9.369 | 10.1 | 9.708 | 10.15 | 9.88 | 10 | 9.97 | 10.2 | 10.6 | > 2 | 10.62 | > 1 | 10.0369 | 11.3781 | 10.1827 | 10.87 | > 2 | > 1 | 9.8316 | 9.3695 | 9.593 | > 1 | 10.24 | > 1 | 9.86 | 9.916 | > 1 | 10.0271 |  | 1
| 1 | 2 | 9.207 | 10 | 9.549 | 10.1 | 9.45 | 9.99 | 10.15 | 10.25 | 10.7 | > 2 | 10.68 | > 1 | 10.0811 | 9.8451 | 10.6482 | 10.85 | > 2 | > 1 | 10.0564 | 9.3838 | 9.445 | 9.89 | 10.25 | > 1 | 9.86 | 10.016 | > 1 | 9.9917 |  | 24
| 1 | 3 | 9.432 | 10.3 | 9.717 | 10.15 | 9.81 | 9.98 | 10.1 | 10.3 | 10.5 | > 2 | 10.64 | > 1 | 10.1138 | 10.7509 | 11.2657 | 10.93 | > 2 | > 1 | 10.0799 | 9.4674 | 9.59 | > 1 | 10.11 | > 1 | 9.78 | 10.072 | > 1 | 10.0081 |  | 16
| 1 | 4 | 9.252 | 10.3 | 9.831 | 10.15 | 9.67 | 10 | 9.93 | 10.9 | 10.7 | > 2 | 10.63 | > 1 | 10.1483 | 10.4368 | 11.0708 | 10.75 | > 2 | > 1 | 10.0512 | 9.2325 | 9.408 | 9.84 | 10.21 | > 1 | 9.79 | 10.019 | > 1 | 10.0185 |  | 10.0724
| 1 | 5 | 9.585 | 10.4 | 9.689 | 10.1 | 9.65 | 10.05 | 10.15 | 10.35 | 10.7 | > 2 | 10.49 | > 1 | 10.1096 | 11.903 | 10.2955 | 10.84 | > 2 | > 1 | 10.0421 | 9.4249 | 9.45 | > 1 | 10.17 | > 1 | 9.81 | 10.067 | > 1 | 10.1145 |  | 117
| 1 | 6 | 9.567 | 10.4 | 9.777 | 10.1 | 9.89 | 10.1 | 10.05 | 10.3 | 10.7 | > 2 | 10.64 | > 1 | 10.154 | 9.1798 | 10.5421 | 10.84 | > 2 | > 1 | 10.0443 | 9.2252 | 9.62 | > 1 | 10.25 | > 1 | 9.84 | 9.884 | > 1 | 10.0795
| Mean |  | 9.402 | 10.25 | 9.7118 | 10.125 | 9.725 | 10.02 | 10.0583 | 10.3833 | 10.65 |  | 10.6167 |  | 10.1073 | 10.5823 | 10.6675 | 10.8467 |  |  | 10.0176 | 9.3506 | 9.5177 | 9.865 | 10.205 |  | 9.8233 | 9.9957 |  | 10.0399
| Median |  | 9.4005 | 10.3 | 9.7125 | 10.125 | 9.74 | 10 | 10.075 | 10.3 | 10.7 |  | 10.635 |  | 10.1117 | 10.5939 | 10.5952 | 10.845 |  |  | 10.0478 | 9.3766 | 9.52 | 9.865 | 10.225 |  | 9.825 | 10.0175 |  | 10.0228
| Std Dev. |  | 0.15699 | 0.16432 | 0.09542 | 0.027386 | 0.16897 | 0.046043 | 0.09261 | 0.2582 | 0.083666 |  | 0.06532 |  | 0.043708 | 0.99377 | 0.4266 | 0.058195 |  |  | 0.092114 | 0.10029 | 0.093022 | 0.035355 | 0.055767 |  | 0.035024 | 0.078345 |  | 0.047095
| Rel.Std.Dev. |  | 1.67% | 1.603% | 0.9825% | 0.2705% | 1.737% | 0.4595% | 0.9207% | 2.487% | 0.7856% |  | 0.6153% |  | 0.4324% | 9.391% | 3.999% | 0.5365% |  |  | 0.9195% | 1.073% | 0.9774% | 0.3584% | 0.5465% |  | 0.3565% | 0.7838% |  | 0.4691%
| PDM3 |  | -6.656% | 1.763% | -3.58% | 0.5221% | -3.449% | -0.5204% | -0.1398% | 3.087% | 5.734% |  | 5.403% |  | 0.346% | 5.062% | 5.908% | 7.687% |  |  | -0.5444% | -7.167% | -5.508% | -2.059% | 1.316% |  | -2.473% | -0.762% |  | -0.3228%
| Z-Score (Absolute) |  | 1.59 | 0.49 | 0.83 | 0.19 | 0.79 | 0.07 | 0.02 | 0.82 | 1.47 | NA | 1.39 | NA | 0.14 | 1.31 | 1.52 | 1.96 | NA | NA | 0.08 | 1.71 | 1.3 | 0.45 | 0.38 | NA | 0.55 | 0.13 | NA | 0.02
| Analytical results for Zr in OREAS 627 (Certified Value 37 ppm) | [30 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.21 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | 1.30 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | [7 empty] | 1
|  |  | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 0.5g | 25g | 0.2g | 01g | 01g | 01g | 0.25g | 0.25g | 0.2g | 0.25g | 0.25g | 0.25g | 0.5g | 0.25g | 0.2g | [7 empty] | 1
| 1 | 1 | 40 | 43.3 | 35.2 | 40.6 | 35.2 | 33.8 | 38.9 | 24.8 | 42.5 | 35.8632 | 34.0111 | 29.4 | 45.3 | 39.7 | 35.14 | 50.1 | 42.5 | 31.2 | 32.8 | 33.9 | 48.2 | 33.2 | [7 empty] | 1
| 1 | 2 | 38 | 44 | 35.1 | 37 | 36 | 36.1 | 38.2 | 25.1 | 41.2 | 35.7032 | 33.5475 | 26.8 | 45.4 | 38.4 | 35.05 | 48.8 | 39.8 | 30.9 | 33.2 | 36.2 | 49.4 | 34.7 | [7 empty] | 15
| 1 | 3 | 40 | 46.1 | 35.1 | 40.9 | 36.2 | 34.4 | 37.7 | 25.3 | 41.1 | 35.2377 | 33.6145 | 28.1 | 46.6 | 38.2 | 34.92 | 49.3 | 40 | 31.5 | 33.9 | 34.6 | 47.6 | 41.9 | [7 empty] | 16
| 1 | 4 | 40 | 43.6 | 35.6 | 39.2 | 32.9 | 36.4 | 38.3 | 24 | 42.7 | 35.8894 | 33.6928 | 24.3 | 45.9 | 37.8 | 35.11 | 48.2 | 40.1 | 31 | 32.5 | 35.3 | 47 | 37.6 | [7 empty] | 37.0008
| 1 | 5 | 40 | 42.5 | 35.8 | 40.6 | 34.1 | 34.6 | 37.1 | 25.6 | 43.2 | 35.9108 | 34.1492 | 28.1 | 46.3 | 38.7 | 35.08 | 47.5 | 39.6 | 31.4 | 31.7 | 34.4 | 48.7 | 40.1 | [7 empty] | 118
| 1 | 6 | 40 | 46 | 35.2 | 40.7 | 35 | 35.1 | 38.3 | 24 | 40.6 | 37.2939 | 34.3156 | 28.7 | 48.1 | 39 | 35.12 | 47.3 | 42.2 | 31.8 | 33.4 | 38.5 | 49.1 | 29.1
| Mean |  | 39.6667 | 44.25 | 35.3333 | 39.8333 | 34.9 | 35.0667 | 38.0833 | 24.8 | 41.8833 | 35.983 | 33.8885 | 27.5667 | 46.2667 | 38.6333 | 35.07 | 48.5333 | 40.7 | 31.3 | 32.9167 | 35.4833 | 48.3333 | 36.1
| Median |  | 40 | 43.8 | 35.2 | 40.6 | 35.1 | 34.85 | 38.25 | 24.95 | 41.85 | 35.8763 | 33.8519 | 28.1 | 46.1 | 38.55 | 35.095 | 48.5 | 40.05 | 31.3 | 33 | 34.95 | 48.45 | 36.15
| Std Dev. |  | 0.8165 | 1.4789 | 0.29439 | 1.5161 | 1.2361 | 1.0113 | 0.61455 | 0.67231 | 1.0496 | 0.69003 | 0.31466 | 1.8151 | 1.0289 | 0.66533 | 0.08 | 1.0783 | 1.2931 | 0.33466 | 0.7679 | 1.6798 | 0.91579 | 4.7172
| Rel.Std.Dev. |  | 2.058% | 3.342% | 0.8332% | 3.806% | 3.542% | 2.884% | 1.614% | 2.711% | 2.506% | 1.918% | 0.9285% | 6.584% | 2.224% | 1.722% | 0.2281% | 2.222% | 3.177% | 1.069% | 2.333% | 4.734% | 1.895% | 13.07%
| PDM3 |  | 7.205% | 19.59% | -4.506% | 7.655% | -5.678% | -5.227% | 2.926% | -32.97% | 13.2% | -2.751% | -8.411% | -25.5% | 25.04% | 4.412% | -5.218% | 31.17% | 9.998% | -15.41% | -11.04% | -4.101% | 30.63% | -2.434%
| Z-Score (Absolute) |  | 0.72 | 1.64 | 0.14 | 0.76 | 0.23 | 0.19 | 0.41 | 2.25 | 1.17 | 0.01 | 0.43 | 1.69 | 2.04 | 0.52 | 0.19 | 2.5 | 0.93 | 0.95 | 0.62 | 0.11 | 2.46 | 0.01
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C162: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C180: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C198: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C216: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C234: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C252: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C270: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C288: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C306: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C324: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C342: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C361: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C379: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C398: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C416: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C435: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C453: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C528: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C546: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C564: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C582: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C600: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C618: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C637: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C655: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C674: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C692: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C710: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C728: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C746: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C764: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C782: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C800: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C818: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C836: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C854: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C872: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C890: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C909: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C928: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C946: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C965: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C984: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1003: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1021: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1113: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1132: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1150: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1168: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## PF ICP
| Analytical results for Ag in OREAS 627 (Indicative Value 63.8 ppm) | [24 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | [19 empty] | 1
| No | No | 1.10 | 1.12 | 1.16 | 1.23 | [19 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*MS | [19 empty] | 1
1
| 1 | 1 | 64 | 65 | 62.64 | 64 | [19 empty] | 1
| 1 | 2 | 63 | 66 | 61.76 | 64 | [19 empty] | 1
| 1 | 3 | 63 | 67 | 62.26 | 65 | [19 empty] | 16
| 1 | 4 | 64 | 65 | 61.25 | 64 | [19 empty] | 63.7625
| 1 | 5 | 62 | 67 | 61.36 | 66 | [19 empty] | 18
| 1 | 6 | 62 | 62 | 62.03 | 66
| Mean |  | 63 | 65.3333 | 61.8833 | 64.8333
| Median |  | 63 | 65.5 | 61.895 | 64.5
| Std Dev. |  | 0.89443 | 1.8619 | 0.53407 | 0.98319
| Rel.Std.Dev. |  | 1.42% | 2.85% | 0.863% | 1.516%
| PDM3 |  | -1.196% | 2.464% | -2.947% | 1.679%
| Z-Score (Absolute) |  | 0.53 | 0.82 | 1.18 | 0.53
| Analytical results for Al in OREAS 627 (Certified Value 5.65 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 2
3
| 1 | 1 | 6.29 | 5.65 | 5.7 | 5.478 | 5.583 | 5.636 | 5.636 | 5.425 | 5.7 | 5.71 | 5.76 | 5.63 | 5.4683 | 5.874 | 5.3 | 5.74 | 5.9464 | 5.7368 | 5.63 | 5.81 | 5.52 | 5.47 |  | 1
| 1 | 2 | 6.4 | 5.63 | 5.83 | 5.531 | 5.663 | 5.663 | 5.61 | 5.451 | 5.75 | 5.56 | 5.57 | 5.49 | 5.5047 | 5.9112 | 5.42 | 5.66 | 6.0309 | 5.7093 | 5.58 | 5.8 | 5.59 | 5.61 |  | 0
| 1 | 3 | 6.37 | 5.67 | 5.73 | 5.531 | 5.636 | 5.689 | 5.689 | 5.531 | 5.79 | 5.65 | 5.69 | 5.58 | 5.5347 | 5.981 | 5.42 | 5.72 | 5.8197 | 5.7187 | 5.66 | 5.85 | 5.54 | 5.53 |  | 16
| 1 | 4 | 6.29 | 5.54 | 5.61 | 5.478 | 5.557 | 5.769 | 5.636 | 5.372 | 5.76 | 5.69 | 5.77 | 5.42 | 5.4686 | 5.8697 | 5.31 | 5.73 | 6.0443 | 5.7395 | 5.61 | 5.89 | 5.49 | 5.46 |  | 5.6503
| 1 | 5 | 6.34 | 5.62 | 5.99 | 5.504 | 5.61 | 5.663 | 5.583 | 5.451 | 5.8 | 5.49 | 5.59 | 5.67 | 5.6088 | 5.9863 | 5.4 | 5.76 | 6.2185 | 5.7325 | 5.68 | 5.87 | 5.48 | 5.68 |  | 120
| 1 | 6 | 6.39 | 5.73 | 5.83 | 5.504 | 5.689 | 5.636 | 5.689 | 5.451 | 5.76 | 5.61 | 5.6 | 5.5 | 5.4762 | 5.953 | 5.37 | 5.71 | 5.8628 | 5.7614 | 5.64 | 5.83 | 5.46 | 5.62
| Mean |  | 6.3467 | 5.64 | 5.7817 | 5.5043 | 5.623 | 5.676 | 5.6405 | 5.4468 | 5.76 | 5.6183 | 5.6633 | 5.5483 | 5.5102 | 5.9292 | 5.37 | 5.72 | 5.9871 | 5.733 | 5.6333 | 5.8417 | 5.5133 | 5.5617
| Median |  | 6.355 | 5.64 | 5.78 | 5.504 | 5.623 | 5.663 | 5.636 | 5.451 | 5.76 | 5.63 | 5.645 | 5.54 | 5.4904 | 5.9321 | 5.385 | 5.725 | 5.9886 | 5.7346 | 5.635 | 5.84 | 5.505 | 5.57
| Std Dev. |  | 0.048442 | 0.06261 | 0.13182 | 0.023704 | 0.049497 | 0.049711 | 0.042383 | 0.051406 | 0.035214 | 0.083046 | 0.088919 | 0.094534 | 0.054754 | 0.051821 | 0.053666 | 0.034059 | 0.14413 | 0.018052 | 0.03559 | 0.034881 | 0.047188 | 0.088863
| Rel.Std.Dev. |  | 0.7633% | 1.11% | 2.28% | 0.4306% | 0.8803% | 0.8758% | 0.7514% | 0.9438% | 0.6113% | 1.478% | 1.57% | 1.704% | 0.9937% | 0.874% | 0.9994% | 0.5954% | 2.407% | 0.3149% | 0.6318% | 0.5971% | 0.8559% | 1.598%
| PDM3 |  | 12.32% | -0.182% | 2.325% | -2.583% | -0.4828% | 0.4552% | -0.1731% | -3.601% | 1.942% | -0.5654% | 0.231% | -1.804% | -2.479% | 4.936% | -4.96% | 1.234% | 5.961% | 1.465% | -0.2999% | 3.387% | -2.424% | -1.568%
| Z-Score (Absolute) |  | 4.49 | 0 | 0.9 | 0.87 | 0.11 | 0.23 | 0 | 1.23 | 0.76 | 0.14 | 0.15 | 0.59 | 0.83 | 1.84 | 1.72 | 0.51 | 2.2 | 0.59 | 0.04 | 1.28 | 0.81 | 0.5
| Analytical results for As in OREAS 627 (Certified Value 315 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  | 0
0
| 1 | 1 | 316 | 332 | 335 | 200 | 300 | 400 | 300 | 400 | 297 | 337 | 300 | 306 | 323.5278 | 275 | 320 | 310 | 340 | 310 | 300 | 300 |  |  |  | 1
| 1 | 2 | 315 | 325 | 316 | 300 | 300 | 500 | 400 | 400 | 318 | 337 | 300 | 295 | 321.6356 | 273.1 | 310 | 313 | 340 | 300 | 300 | 300 |  |  |  | 21
| 1 | 3 | 313 | 309 | 306 | 300 | 400 | 400 | 400 | 400 | 312 | 337 | 400 | 296 | 350.0073 | 269 | 320 | 321 | 341 | 290 | 300 | 200 |  |  |  | 16
| 1 | 4 | 310 | 327 | 313 | 300 | 300 | 400 | 300 | 300 | 320 | 348 | 400 | 298 | 355.4624 | 278.8 | 320 | 306 | 341 | 300 | 300 | 300 |  |  |  | 315.2944
| 1 | 5 | 315 | 334 | 316 | 300 | 300 | 400 | 400 | 400 | 297 | 337 | 400 | 306 | 342.479 | 277.2 | 330 | 318 | 341 | 300 | 300 | 200 |  |  |  | 121
| 1 | 6 | 318 | 328 | 300 | 200 | 300 | 400 | 400 | 400 | 325 | 292 | 300 | 311 | 330.2495 | 281.1 | 320 | 311 | 339 | 310 | 300 | 300
| Mean |  | 314.5 | 325.8333 | 314.3333 | 266.6667 | 316.6667 | 416.6667 | 366.6667 | 383.3333 | 311.5 | 331.3333 | 350 | 302 | 337.2269 | 275.7 | 320 | 313.1667 | 340.3333 | 301.6667 | 300 | 266.6667
| Median |  | 315 | 327.5 | 314.5 | 300 | 300 | 400 | 400 | 400 | 315 | 337 | 350 | 302 | 336.3643 | 276.1 | 320 | 312 | 340.5 | 300 | 300 | 300
| Std Dev. |  | 2.7386 | 8.8863 | 11.9108 | 51.6398 | 40.8248 | 40.8248 | 51.6398 | 40.8248 | 11.9791 | 19.7653 | 54.7723 | 6.5422 | 14.1626 | 4.3165 | 6.3246 | 5.4924 | 0.8165 | 7.5277 | 0 | 51.6398
| Rel.Std.Dev. |  | 0.8708% | 2.727% | 3.789% | 19.36% | 12.89% | 9.798% | 14.08% | 10.65% | 3.846% | 5.965% | 15.65% | 2.166% | 4.2% | 1.566% | 1.976% | 1.754% | 0.2399% | 2.495% | 0% | 19.36%
| PDM3 |  | -0.2519% | 3.343% | -0.3048% | -15.42% | 0.4352% | 32.15% | 16.29% | 21.58% | -1.203% | 5.087% | 11.01% | -4.216% | 6.956% | -12.56% | 1.492% | -0.6748% | 7.941% | -4.322% | -4.851% | -15.42%
| Z-Score (Absolute) |  | 0 | 0.65 | 0 | 2.7 | 0.13 | NA | NA | NA | 0.17 | 0.96 | NA | 0.7 | 1.29 | 2.19 | 0.32 | 0.07 | 1.47 | 0.72 | NA | NA
Results from laboratories 1.06, 1.07, 1.08, 1.11, 1.27 and 1.28 were removed due to their 100 ppm reading resolution.
| Analytical results for B in OREAS 627 (Indicative Value 26.8 ppm) | [24 empty] | Indicative
B | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.20; 1.06: 1.26 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES; 1.05: PF*MS; 1.06: PF*OES |  |  |  | 1
1
| 1 | 1 | 1.01: 20; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 32; 1.06: 30 |  |  |  | 1
| 1 | 2 | 1.01: 20; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 34; 1.06: 20 |  |  |  | 3
| 1 | 3 | 1.01: 10; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 34; 1.06: 30 |  |  |  | 16
| 1 | 4 | 1.01: 20; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 40; 1.06: 30 |  |  |  | 26.7556
| 1 | 5 | 1.01: 20; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 34; 1.06: 30 |  |  |  | 19
| 1 | 6 | 1.01: 20; 1.02: < 20; 1.03: < 50; 1.04: < 50; 1.05: 34; 1.06: 30
| Mean |  | 1.01: 18.3333; 1.05: 34.6667; 1.06: 28.3333
| Median |  | 1.01: 20; 1.05: 34; 1.06: 30
| Std Dev. |  | 1.01: 4.0825; 1.05: 2.7325; 1.06: 4.0825
| Rel.Std.Dev. |  | 1.01: 22.27%; 1.05: 7.882%; 1.06: 14.41%
| PDM3 |  | 1.01: -31.48%; 1.05: 29.57%; 1.06: 5.897%
| Z-Score (Absolute) |  | 1.01: 0.9; 1.02: 2.02; 1.03: 0; 1.04: 0; 1.05: 1.3; 1.06: 0.45
| Analytical results for Ba in OREAS 627 (Certified Value 4384 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | [9 empty] | 0
0
| 1 | 1 | 4667 | 4310 | 4236 | 4450 | 4394 | 4250 | 4271.2355 | 4579.4351 | 4470.8 | 4720 | 4729 | 4207 | 4137 | 3890 | [9 empty] | 1
| 1 | 2 | 4665 | 4390 | 4304 | 4510 | 4426 | 4176 | 4107.2657 | 4550.7404 | 4480 | 4740 | 4786 | 4183 | 4114 | 3870 | [9 empty] | 4
| 1 | 3 | 4622 | 4340 | 4232 | 4520 | 4394 | 4214 | 4079.3305 | 4581.3341 | 4552.2 | 4710 | 4420 | 4195 | 4152 | 3920 | [9 empty] | 16
| 1 | 4 | 4630 | 4370 | 4149 | 4490 | 4475 | 4111 | 4170.6817 | 4551.5669 | 4471.2 | 4830 | 4801 | 4112 | 4077 | 3900 | [9 empty] | 4383.897
| 1 | 5 | 4609 | 4360 | 4478 | 4520 | 4424 | 4320 | 4214.5278 | 4535.1274 | 4489.3 | 4960 | 4917 | 4171 | 4138 | 3860 | [9 empty] | 122
| 1 | 6 | 4689 | 4360 | 4382 | 4480 | 4325 | 4167 | 4160.533 | 4555.4669 | 4377.8 | 4790 | 4736 | 4197 | 4094 | 3880
| Mean |  | 4647 | 4355 | 4296.8333 | 4495 | 4406.3333 | 4206.3333 | 4167.2624 | 4558.9451 | 4473.55 | 4791.6667 | 4731.5 | 4177.5 | 4118.6667 | 3886.6667
| Median |  | 4647.5 | 4360 | 4270 | 4500 | 4409 | 4195 | 4165.6074 | 4553.5169 | 4475.6 | 4765 | 4761 | 4189 | 4125.5 | 3885
| Std Dev. |  | 31.132 | 27.3861 | 118.2242 | 27.3861 | 49.6655 | 72.7205 | 69.9042 | 18.0125 | 55.9583 | 94.1099 | 166.8949 | 34.4078 | 28.939 | 21.6025
| Rel.Std.Dev. |  | 0.6699% | 0.6288% | 2.751% | 0.6093% | 1.127% | 1.729% | 1.677% | 0.3951% | 1.251% | 1.964% | 3.527% | 0.8236% | 0.7026% | 0.5558%
| PDM3 |  | 6.002% | -0.6592% | -1.986% | 2.534% | 0.5118% | -4.05% | -4.942% | 3.993% | 2.045% | 9.302% | 7.929% | -4.708% | -6.05% | -11.34%
| Z-Score (Absolute) |  | 0.94 | 0.09 | 0.3 | 0.4 | 0.09 | 0.62 | 0.75 | 0.63 | 0.33 | 1.45 | 1.24 | 0.72 | 0.93 | 1.75
| Analytical results for Be in OREAS 627 (Indicative Value 1.78 ppm) | [24 empty] | Indicative
Be | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.16; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23; 1.09: 1.25; 1.10: 1.26 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES; 1.09: PF*OES; 1.10: PF*OES |  |  |  | 2
2
| 1 | 1 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 2.4305; 1.05: < 2; 1.06: 2; 1.07: 0.3; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 1
| 1 | 2 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 2.7432; 1.05: < 2; 1.06: 2; 1.07: 0.7; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 14
| 1 | 3 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 2.2474; 1.05: < 2; 1.06: 2; 1.07: 1.2; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 16
| 1 | 4 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 2.0077; 1.05: < 2; 1.06: 2; 1.07: 0.7; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 1.7838
| 1 | 5 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 2.1422; 1.05: < 2; 1.06: 2; 1.07: 1.1; 1.08: < 5; 1.09: < 5; 1.10: < 5 |  |  |  | 20
| 1 | 6 | 1.01: < 3; 1.02: < 5; 1.03: 2; 1.04: 2.7408; 1.05: < 2; 1.06: 2; 1.07: 0.5; 1.08: < 5; 1.09: < 5; 1.10: < 5
| Mean |  | 1.03: 2; 1.04: 2.3853; 1.06: 2; 1.07: 0.75
| Median |  | 1.03: 2; 1.04: 2.3389; 1.06: 2; 1.07: 0.7
| Std Dev. |  | 1.03: 0; 1.04: 0.30894; 1.06: 0; 1.07: 0.34496
| Rel.Std.Dev. |  | 1.03: 0%; 1.04: 12.95%; 1.06: 0%; 1.07: 46%
| PDM3 |  | 1.03: 12.12%; 1.04: 33.72%; 1.06: 12.12%; 1.07: -57.96%
| Z-Score (Absolute) |  | 1.01: 1.52; 1.02: 0.67; 1.03: 0.42; 1.04: 0.42; 1.05: 2.62; 1.06: 0.42; 1.07: 3.17; 1.08: 0.67; 1.09: 0.67; 1.10: 0.67
| Analytical results for Bi in OREAS 627 (Certified Value 43.7 ppm) | [24 empty] | Certified
Bi | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.16; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*MS |  |  |  | 1
2
| 1 | 1 | 1.01: 39; 1.02: 39.4; 1.03: 45.8; 1.04: 48.5759; 1.05: 44.7; 1.06: 485000; 1.07: 42; 1.08: 45.3 |  |  |  | 1
| 1 | 2 | 1.01: 39; 1.02: 41.1; 1.03: 46.4; 1.04: 47.7027; 1.05: 43.8; 1.06: 471000; 1.07: 42; 1.08: 44.6 |  |  |  | 35
| 1 | 3 | 1.01: 38; 1.02: 41.7; 1.03: 46.2; 1.04: 48.9014; 1.05: 44.5; 1.06: 476000; 1.07: 40; 1.08: 46.2 |  |  |  | 16
| 1 | 4 | 1.01: 38; 1.02: 42.8; 1.03: 46.7; 1.04: 48.7774; 1.05: 45.1; 1.06: 471000; 1.07: 42; 1.08: 45.6 |  |  |  | 43.6962
| 1 | 5 | 1.01: 39; 1.02: 40.3; 1.03: 45.6; 1.04: 47.7178; 1.05: 43.8; 1.06: 477000; 1.07: 41; 1.08: 45.9 |  |  |  | 123
| 1 | 6 | 1.01: 38; 1.02: 42.3; 1.03: 46.3; 1.04: 48.0637; 1.05: 43.9; 1.06: 481000; 1.07: 42; 1.08: 47.5
| Mean |  | 1.01: 38.5; 1.02: 41.2667; 1.03: 46.1667; 1.04: 48.2898; 1.05: 44.3; 1.06: 476833.3333; 1.07: 41.5; 1.08: 45.85
| Median |  | 1.01: 38.5; 1.02: 41.4; 1.03: 46.25; 1.04: 48.3198; 1.05: 44.2; 1.06: 476500; 1.07: 42; 1.08: 45.75
| Std Dev. |  | 1.01: 0.54772; 1.02: 1.2691; 1.03: 0.40332; 1.04: 0.53229; 1.05: 0.54772; 1.06: 5528.7129; 1.07: 0.83666; 1.08: 0.97724
| Rel.Std.Dev. |  | 1.01: 1.423%; 1.02: 3.075%; 1.03: 0.8736%; 1.04: 1.102%; 1.05: 1.236%; 1.06: 1.159%; 1.07: 2.016%; 1.08: 2.131%
| PDM3 |  | 1.01: -11.89%; 1.02: -5.56%; 1.03: 5.654%; 1.04: 10.51%; 1.05: 1.382%; 1.06: 1091148%; 1.07: -5.026%; 1.08: 4.929%
| Z-Score (Absolute) |  | 1.01: 1.31; 1.02: 0.76; 1.03: 0.22; 1.04: 0.64; 1.05: 0.15; 1.06: 94701.33; 1.07: 0.71; 1.08: 0.15
| Analytical results for Ca in OREAS 627 (Certified Value 0.782 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  |  | 3
3
| 1 | 1 | 0.78 | 0.79 | 0.84 | 0.636 | 0.886 | 0.729 | 0.758 | 0.786 | 0.77 | 0.8 | 0.74 | 0.8 | 0.83914 | 0.78 | 0.77 | 1.0094 | 0.8 | 0.82 | 0.7 | 0.79 |  |  |  | 1
| 1 | 2 | 0.79 | 0.77 | 0.96 | 0.65 | 0.772 | 0.75 | 0.779 | 0.779 | 0.76 | 0.8 | 0.75 | 0.8 | 0.83956 | 0.78 | 0.75 | 0.7359 | 0.8 | 0.81 | 0.69 | 0.8 |  |  |  | 0
| 1 | 3 | 0.78 | 0.81 | 0.93 | 0.658 | 0.815 | 0.765 | 0.786 | 0.779 | 0.75 | 0.8 | 0.79 | 0.8 | 0.86383 | 0.82 | 0.76 | 0.7241 | 0.8 | 0.81 | 0.7 | 0.78 |  |  |  | 16
| 1 | 4 | 0.77 | 0.75 | 0.92 | 0.593 | 0.815 | 0.765 | 0.8 | 0.629 | 0.75 | 0.9 | 0.8 | 0.8 | 0.85557 | 0.76 | 0.76 | 0.78 | 0.8 | 0.82 | 0.69 | 0.77 |  |  |  | 0.78168
| 1 | 5 | 0.79 | 0.78 | 0.92 | 0.643 | 0.772 | 0.736 | 0.772 | 0.772 | 0.77 | 0.8 | 0.75 | 0.8 | 0.84866 | 0.8 | 0.76 | 0.8167 | 0.8 | 0.85 | 0.69 | 0.77 |  |  |  | 124
| 1 | 6 | 0.78 |  | 0.74 | 0.629 | 0.829 | 0.729 | 0.786 | 0.779 | 0.75 | 0.7 | 0.74 | 0.8 | 0.83055 | 0.76 | 0.76 | 0.7365 | 0.8 | 0.84 | 0.69 | 0.79
| Mean |  | 0.78167 | 0.78 | 0.885 | 0.63483 | 0.81483 | 0.74567 | 0.78017 | 0.754 | 0.75833 | 0.8 | 0.76167 | 0.8 | 0.84622 | 0.78333 | 0.76 | 0.80043 | 0.8 | 0.825 | 0.69333 | 0.78333
| Median |  | 0.78 | 0.78 | 0.92 | 0.6395 | 0.815 | 0.743 | 0.7825 | 0.779 | 0.755 | 0.8 | 0.75 | 0.8 | 0.84411 | 0.78 | 0.76 | 0.75825 | 0.8 | 0.82 | 0.69 | 0.785
| Std Dev. |  | 0.0075277 | 0.022361 | 0.081425 | 0.022886 | 0.042273 | 0.016825 | 0.014289 | 0.061397 | 0.0098319 | 0.063246 | 0.026394 | 1.2162e-16 | 0.012194 | 0.023381 | 0.0063246 | 0.1081 | 1.2162e-16 | 0.016432 | 0.005164 | 0.012111
| Rel.Std.Dev. |  | 0.963% | 2.867% | 9.201% | 3.605% | 5.188% | 2.256% | 1.831% | 8.143% | 1.297% | 7.906% | 3.465% | 1.52e-14% | 1.441% | 2.985% | 0.8322% | 13.51% | 1.52e-14% | 1.992% | 0.7448% | 1.546%
| PDM3 |  | -0.001858% | -0.2151% | 13.22% | -18.79% | 4.241% | -4.607% | -0.1938% | -3.541% | -2.987% | 2.344% | -2.56% | 2.344% | 8.256% | 0.2114% | -2.774% | 2.399% | 2.344% | 5.542% | -11.3% | 0.2114%
| Z-Score (Absolute) |  | 0.04 | 0.01 | 3.23 | 4.48 | 1.06 | 1.06 | 0 | 0.81 | 0.67 | NA | 0.57 | NA | 2.03 | 0.1 | 0.62 | 0.62 | NA | 1.38 | 2.68 | 0.1
Results from laboratories 1.10, 1.12 and 1.23 were removed due to their 0.1 wt.% reading resolution.
| Analytical results for Cd in OREAS 627 (Certified Value 263 ppm) | [24 empty] | Certified
Cd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab |  |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.09; 1.05: 1.10; 1.06: 1.15; 1.07: 1.16; 1.08: 1.23; 1.09: 1.25; 1.10: 1.26 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*OES; 1.10: PF*OES |  |  |  | 0
0
| 1 | 1 | 1.01: 262; 1.02: 277; 1.03: 239; 1.04: 243; 1.05: 271; 1.06: 245.399; 1.07: 262.5; 1.08: 259; 1.09: 259; 1.10: 220 |  |  |  | 1
| 1 | 2 | 1.01: 266; 1.02: 285; 1.03: 243; 1.04: 263; 1.05: 272; 1.06: 247.3408; 1.07: 260.1; 1.08: 263; 1.09: 255; 1.10: 210 |  |  |  | 0
| 1 | 3 | 1.01: 264; 1.02: 276; 1.03: 234; 1.04: 264; 1.05: 270; 1.06: 253.906; 1.07: 257.4; 1.08: 266; 1.09: 256; 1.10: 210 |  |  |  | 16
| 1 | 4 | 1.01: 262; 1.02: 275; 1.03: 236; 1.04: 266; 1.05: 273; 1.06: 252.3678; 1.07: 263.8; 1.08: 260; 1.09: 257; 1.10: 210 |  |  |  | 262.9101
| 1 | 5 | 1.01: 263; 1.02: 268; 1.03: 239; 1.04: 252; 1.05: 265; 1.06: 251.3154; 1.07: 262.6; 1.08: 273; 1.09: 254; 1.10: 210 |  |  |  | 125
| 1 | 6 | 1.01: 262; 1.02: 277; 1.03: 234; 1.04: 267; 1.05: 263; 1.06: 242.2576; 1.07: 264.3; 1.08: 272; 1.09: 257; 1.10: 210
| Mean |  | 1.01: 263.1667; 1.02: 276.3333; 1.03: 237.5; 1.04: 259.1667; 1.05: 269; 1.06: 248.7644; 1.07: 261.7833; 1.08: 265.5; 1.09: 256.3333; 1.10: 211.6667
| Median |  | 1.01: 262.5; 1.02: 276.5; 1.03: 237.5; 1.04: 263.5; 1.05: 270.5; 1.06: 249.3281; 1.07: 262.55; 1.08: 264.5; 1.09: 256.5; 1.10: 210
| Std Dev. |  | 1.01: 1.6021; 1.02: 5.4283; 1.03: 3.5071; 1.04: 9.5795; 1.05: 4.0497; 1.06: 4.5082; 1.07: 2.5934; 1.08: 5.9582; 1.09: 1.7512; 1.10: 4.0825
| Rel.Std.Dev. |  | 1.01: 0.6088%; 1.02: 1.964%; 1.03: 1.477%; 1.04: 3.696%; 1.05: 1.505%; 1.06: 1.812%; 1.07: 0.9907%; 1.08: 2.244%; 1.09: 0.6832%; 1.10: 1.929%
| PDM3 |  | 1.01: 0.09757%; 1.02: 5.106%; 1.03: -9.665%; 1.04: -1.424%; 1.05: 2.316%; 1.06: -5.38%; 1.07: -0.4286%; 1.08: 0.9851%; 1.09: -2.502%; 1.10: -19.49%
| Z-Score (Absolute) |  | 1.01: 0.27; 1.02: 1.58; 1.03: 2.29; 1.04: 0.13; 1.05: 0.85; 1.06: 1.17; 1.07: 0.13; 1.08: 0.5; 1.09: 0.41; 1.10: 4.86
| Analytical results for Ce in OREAS 627 (Certified Value 70 ppm) | [24 empty] | Certified
Ce | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.23 |  |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS |  |  |  | 0
1
| 1 | 1 | 1.01: 69.9; 1.02: 61.6; 1.03: 73; 1.04: 70.5124; 1.05: 75.418; 1.06: 109; 1.07: 68.5 |  |  |  | 1
| 1 | 2 | 1.01: 67.5; 1.02: 64.4; 1.03: 71; 1.04: 67.8347; 1.05: 75.614; 1.06: 75; 1.07: 66.2 |  |  |  | 36
| 1 | 3 | 1.01: 69.1; 1.02: 64.7; 1.03: 73.1; 1.04: 68.4742; 1.05: 75.75; 1.06: 71; 1.07: 70 |  |  |  | 16
| 1 | 4 | 1.01: 60; 1.02: 67.2; 1.03: 73.1; 1.04: 70.455; 1.05: 75.635; 1.06: 80; 1.07: 70.6 |  |  |  | 70.4741
| 1 | 5 | 1.01: 71.8; 1.02: 61.9; 1.03: 71.3; 1.04: 71.459; 1.05: 75.597; 1.06: 86; 1.07: 72.3 |  |  |  | 126
| 1 | 6 | 1.01: 65.4; 1.02: 66.4; 1.03: 69.9; 1.04: 69.6172; 1.05: 75.745; 1.06: 73; 1.07: 70.4
| Mean |  | 1.01: 67.2833; 1.02: 64.3667; 1.03: 71.9; 1.04: 69.7254; 1.05: 75.6265; 1.06: 82.3333; 1.07: 69.6667
| Median |  | 1.01: 68.3; 1.02: 64.55; 1.03: 72.15; 1.04: 70.0361; 1.05: 75.6245; 1.06: 77.5; 1.07: 70.2
| Std Dev. |  | 1.01: 4.1758; 1.02: 2.2809; 1.03: 1.3609; 1.04: 1.3645; 1.05: 0.12154; 1.06: 14.1374; 1.07: 2.0897
| Rel.Std.Dev. |  | 1.01: 6.206%; 1.02: 3.544%; 1.03: 1.893%; 1.04: 1.957%; 1.05: 0.1607%; 1.06: 17.17%; 1.07: 3%
| PDM3 |  | 1.01: -4.528%; 1.02: -8.666%; 1.03: 2.023%; 1.04: -1.062%; 1.05: 7.311%; 1.06: 16.83%; 1.07: -1.146%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 1.48; 1.03: 0.6; 1.04: 0; 1.05: 1.63; 1.06: 3.48; 1.07: 0.02
| Analytical results for Co in OREAS 627 (Certified Value 8.87 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES |  |  | 2
2
| 1 | 1 | 10 | 8.1 | 11 | 20 | < 20 | < 20 | < 20 | < 20 | 7 | 8 | < 20 | < 20 | < 10 | 8.2142 | 7.7 | 9 | 15 | 7.6 | < 10 | 10 | < 20 |  |  | 1
| 1 | 2 | 11 | 8.2 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 8 | 8 | < 20 | < 20 | < 10 | 8.0924 | 7.7 | 9 | 9 | 7.8 | < 10 | 10 | < 20 |  |  | 37
| 1 | 3 | 11 | 7.8 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 8 | 8 | < 20 | < 20 | < 10 | 7.7632 | 7.4 | 9 | 10 | 7.9 | < 10 | 10 | < 20 |  |  | 16
| 1 | 4 | 11 | 7.9 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 8 | 8 | < 20 | < 20 | < 10 | 7.6803 | 7.7 | 9 | 10 | 7.5 | < 10 | 10 | < 20 |  |  | 8.8723
| 1 | 5 | 11 | 8.2 | 11 | < 20 | < 20 | < 20 | < 20 | < 20 | 7 | 8 | < 20 | < 20 | < 10 | 7.9498 | 7.9 | 9 | 9 | 8 | < 10 | 10 | < 20 |  |  | 127
| 1 | 6 | 11 | 7.7 | < 10 | < 20 | < 20 | < 20 | < 20 | < 20 | 8 | 8 | < 20 | < 20 | < 10 | 7.5722 | 7.6 | 10 | 10 | 8 | < 10 | 10 | < 20
| Mean |  | 1.01: 10.8333; 1.02: 7.9833; 1.03: 11; 1.04: 20; 1.09: 7.6667; 1.10: 8; 1.15: 7.8787; 1.16: 7.6667; 1.17: 9.1667; 1.20: 10.5; 1.23: 7.8; 1.26: 10
| Median |  | 1.01: 11; 1.02: 8; 1.03: 11; 1.04: 20; 1.09: 8; 1.10: 8; 1.15: 7.8565; 1.16: 7.7; 1.17: 9; 1.20: 10; 1.23: 7.85; 1.26: 10
| Std Dev. |  | 1.01: 0.40825; 1.02: 0.2137; 1.03: 0; 1.09: 0.5164; 1.10: 0; 1.15: 0.24899; 1.16: 0.1633; 1.17: 0.40825; 1.20: 2.2583; 1.23: 0.20976; 1.26: 0
| Rel.Std.Dev. |  | 1.01: 3.768%; 1.02: 2.677%; 1.03: 0%; 1.09: 6.736%; 1.10: 0%; 1.15: 3.16%; 1.16: 2.13%; 1.17: 4.454%; 1.20: 21.51%; 1.23: 2.689%; 1.26: 0%
| PDM3 |  | 1.01: 22.1%; 1.02: -10.02%; 1.03: 23.98%; 1.04: 125.4%; 1.09: -13.59%; 1.10: -9.832%; 1.15: -11.2%; 1.16: -13.59%; 1.17: 3.318%; 1.20: 18.35%; 1.23: -12.09%; 1.26: 12.71%
| Z-Score (Absolute) |  | 0.67 | 1.63 | 2.43 | 1.35 | 0 | 0 | 0 | 0 | 1.89 | 1.62 | 0 | 0 | 4.05 | 1.72 | 1.89 | 0.67 | 0.4 | 1.78 | 4.05 | 0 | 0
| Analytical results for Cr in OREAS 627 (Indicative Value 55 ppm) | [24 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.11 | 1.12 | 1.15 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [6 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [6 empty] | 0
0
| 1 | 1 | 60 | 118 | < 70 | 70 | 70 | 70 | 70 | 20 | < 100 | < 50 | < 50 | < 50 | 22 | 16 | 22 | 30 | < 100 | [6 empty] | 1
| 1 | 2 | 60 | < 100 | < 70 | 70 | 70 | 70 | 70 | 20 | < 100 | < 50 | < 50 | < 50 | 22 | 19 | 22 | 30 | < 100 | [6 empty] | 15
| 1 | 3 | 50 | < 100 | < 70 | 70 | 70 | 210 | 70 | < 20 | < 100 | < 50 | < 50 | < 50 | 21 | 13 | 22 | 30 | < 100 | [6 empty] | 16
| 1 | 4 | 50 | 103 | < 70 | 70 | 70 | 70 | 70 | 20 | < 100 | < 50 | < 50 | < 50 | 24 | 16 | 22 | 40 | < 100 | [6 empty] | 55.0184
| 1 | 5 | 50 | 181 | 70 | 70 | 70 | 140 | 70 | 20 | < 100 | < 50 | < 50 | < 50 | 23 | 12 | 23 | 30 | < 100 | [6 empty] | 21
| 1 | 6 | 80 | < 100 | < 70 | 140 | 70 | 140 | 70 | 20 | < 100 | < 50 | < 50 | < 50 | 20 | 11 | 23 | 30 | < 100
| Mean |  | 1.01: 58.3333; 1.02: 134; 1.03: 70; 1.04: 81.6667; 1.05: 70; 1.06: 116.6667; 1.07: 70; 1.08: 20; 1.13: 22; 1.15: 14.5; 1.16: 22.3333; 1.17: 31.6667
| Median |  | 1.01: 55; 1.02: 118; 1.03: 70; 1.04: 70; 1.05: 70; 1.06: 105; 1.07: 70; 1.08: 20; 1.13: 22; 1.15: 14.5; 1.16: 22; 1.17: 30
| Std Dev. |  | 1.01: 11.6905; 1.02: 41.3884; 1.04: 28.5774; 1.05: 0; 1.06: 57.1548; 1.07: 0; 1.08: 0; 1.13: 1.4142; 1.15: 3.0166; 1.16: 0.5164; 1.17: 4.0825
| Rel.Std.Dev. |  | 1.01: 20.04%; 1.02: 30.89%; 1.04: 34.99%; 1.05: 0%; 1.06: 48.99%; 1.07: 0%; 1.08: 0%; 1.13: 6.428%; 1.15: 20.8%; 1.16: 2.312%; 1.17: 12.89%
| PDM3 |  | 1.01: 6.025%; 1.02: 143.6%; 1.03: 27.23%; 1.04: 48.44%; 1.05: 27.23%; 1.06: 112.1%; 1.07: 27.23%; 1.08: -63.65%; 1.13: -60.01%; 1.15: -73.65%; 1.16: -59.41%; 1.17: -42.44%
| Z-Score (Absolute) |  | 0.69 | 1.96 | 0 | 1.5 | 1.07 | 2.79 | 1.07 | 0.81 | 0.38 | 0.56 | 0.56 | 0.56 | 0.67 | 0.96 | 0.66 | 0.31 | 0.38
| Analytical results for Cs in OREAS 627 (Certified Value 5.14 ppm) | [24 empty] | Certified
Cs | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.15; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 5.1; 1.02: 4; 1.03: 5.5; 1.04: 4.9032; 1.05: 5.4755; 1.06: 5.6; 1.07: 6; 1.08: 4.9 |  |  | 1
| 1 | 2 | 1.01: 5; 1.02: 4.5; 1.03: 5.5; 1.04: 4.9603; 1.05: 5.3647; 1.06: 5.8; 1.07: 6; 1.08: 4.9 |  |  | 39
| 1 | 3 | 1.01: 5.4; 1.02: 4.6; 1.03: 5.3; 1.04: 4.9885; 1.05: 5.3348; 1.06: 5.7; 1.07: 6; 1.08: 5 |  |  | 16
| 1 | 4 | 1.01: 5.2; 1.02: 4.7; 1.03: 4.9; 1.04: 5.2404; 1.05: 5.0473; 1.06: 5.9; 1.07: 6; 1.08: 4.9 |  |  | 5.1434
| 1 | 5 | 1.01: 5.2; 1.02: 4.3; 1.03: 5.1; 1.04: 5.2991; 1.05: 5.6708; 1.06: 5.6; 1.07: 6; 1.08: 5.1 |  |  | 128
| 1 | 6 | 1.01: 5.3; 1.02: 4.6; 1.03: 4.7; 1.04: 5.0373; 1.05: 4.8795; 1.06: 5.7; 1.07: 6; 1.08: 5.1
| Mean |  | 1.01: 5.2; 1.02: 4.45; 1.03: 5.1667; 1.04: 5.0715; 1.05: 5.2954; 1.06: 5.7167; 1.07: 6; 1.08: 4.9833
| Median |  | 1.01: 5.2; 1.02: 4.55; 1.03: 5.2; 1.04: 5.0129; 1.05: 5.3497; 1.06: 5.7; 1.07: 6; 1.08: 4.95
| Std Dev. |  | 1.01: 0.14142; 1.02: 0.25884; 1.03: 0.3266; 1.04: 0.16068; 1.05: 0.28783; 1.06: 0.1169; 1.07: 0; 1.08: 0.098319
| Rel.Std.Dev. |  | 1.01: 2.72%; 1.02: 5.817%; 1.03: 6.321%; 1.04: 3.168%; 1.05: 5.435%; 1.06: 2.045%; 1.07: 0%; 1.08: 1.973%
| PDM3 |  | 1.01: 1.101%; 1.02: -13.48%; 1.03: 0.4528%; 1.04: -1.398%; 1.05: 2.956%; 1.06: 11.15%; 1.07: 16.65%; 1.08: -3.112%
| Z-Score (Absolute) |  | 1.01: 0.17; 1.02: 3.75; 1.03: 0; 1.04: 0.5; 1.05: 0.67; 1.06: 2.88; 1.07: NA; 1.08: 0.96
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Cu in OREAS 627 (Certified Value 0.482 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.4928 | 0.513 | 0.481 | 0.456 | 0.478 | 0.481 | 0.467 | 0.465 | 0.47 | 0.4857 | 0.477 | 0.4811 | 0.52434 | 0.55911 | 0.477 | 0.3567 | 0.4904 | 0.457 | 0.4969 | 0.472 | 0.466 | 0.48 |  | 1
| 1 | 2 | 0.4947 | 0.511 | 0.495 | 0.457 | 0.484 | 0.485 | 0.46 | 0.475 | 0.475 | 0.4782 | 0.484 | 0.4741 | 0.5268 | 0.55725 | 0.477 | 0.3544 | 0.4991 | 0.4602 | 0.4949 | 0.487 | 0.479 | 0.48 |  | 25
| 1 | 3 | 0.4904 | 0.5 | 0.4932 | 0.463 | 0.479 | 0.482 | 0.461 | 0.473 | 0.478 | 0.4711 | 0.483 | 0.477 | 0.5303 | 0.56344 | 0.475 | 0.3574 | 0.5099 | 0.4626 | 0.4906 | 0.484 | 0.473 | 0.48 |  | 16
| 1 | 4 | 0.4938 | 0.503 | 0.4794 | 0.465 | 0.475 | 0.494 | 0.467 | 0.46 | 0.473 | 0.4807 | 0.5 | 0.4624 | 0.52487 | 0.56343 | 0.477 | 0.3567 | 0.4933 | 0.459 | 0.4956 | 0.48 | 0.484 | 0.48 |  | 0.48222
| 1 | 5 | 0.49 | 0.5 | 0.5199 | 0.459 | 0.478 | 0.487 | 0.459 | 0.466 | 0.481 | 0.4645 | 0.47 | 0.4848 | 0.51337 | 0.56177 | 0.475 | 0.3578 | 0.501 | 0.4625 | 0.4987 | 0.485 | 0.474 | 0.48 |  | 129
| 1 | 6 | 0.4914 | 0.512 | 0.4962 | 0.456 | 0.483 | 0.484 | 0.464 | 0.471 | 0.475 | 0.4761 | 0.477 | 0.4704 | 0.51351 | 0.55879 | 0.476 | 0.3594 | 0.4905 | 0.4631 | 0.4956 | 0.481 | 0.47 | 0.48
| Mean |  | 0.49218 | 0.5065 | 0.49412 | 0.45933 | 0.4795 | 0.4855 | 0.463 | 0.46833 | 0.47533 | 0.47605 | 0.48183 | 0.47497 | 0.5222 | 0.56063 | 0.47617 | 0.35707 | 0.49737 | 0.46073 | 0.49538 | 0.4815 | 0.47433 | 0.48
| Median |  | 0.4921 | 0.507 | 0.4941 | 0.458 | 0.4785 | 0.4845 | 0.4625 | 0.4685 | 0.475 | 0.47715 | 0.48 | 0.47555 | 0.5246 | 0.56044 | 0.4765 | 0.35705 | 0.4962 | 0.46135 | 0.4956 | 0.4825 | 0.4735 | 0.48
| Std Dev. |  | 0.0018915 | 0.0061563 | 0.014556 | 0.0038297 | 0.0033912 | 0.0046797 | 0.0035214 | 0.0056451 | 0.0038297 | 0.0074468 | 0.010226 | 0.0079746 | 0.0071003 | 0.0026146 | 0.00098319 | 0.0016415 | 0.0075566 | 0.0024262 | 0.0027022 | 0.0053198 | 0.0064083 | 0
| Rel.Std.Dev. |  | 0.3843% | 1.215% | 2.946% | 0.8338% | 0.7072% | 0.9639% | 0.7606% | 1.205% | 0.8057% | 1.564% | 2.122% | 1.679% | 1.36% | 0.4664% | 0.2065% | 0.4597% | 1.519% | 0.5266% | 0.5455% | 1.105% | 1.351% | 0%
| PDM3 |  | 2.067% | 5.036% | 2.468% | -4.745% | -0.5633% | 0.6809% | -3.985% | -2.879% | -1.427% | -1.279% | -0.07947% | -1.503% | 8.291% | 16.26% | -1.255% | -25.95% | 3.142% | -4.455% | 2.73% | -0.1486% | -1.635% | -0.4597%
| Z-Score (Absolute) |  | 0.7 | 1.51 | 0.81 | 1.15 | 0.01 | 0.33 | 0.95 | 0.65 | 0.25 | 0.21 | 0.12 | 0.27 | 2.4 | 4.57 | 0.2 | 6.94 | 1 | 1.08 | 0.88 | 0.1 | 0.31 | 0.01
| Analytical results for Dy in OREAS 627 (Certified Value 3.42 ppm) | [24 empty] | Certified
Dy | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 2
3
| 1 | 1 | 1.01: 3.1; 1.02: 3.2; 1.03: 3.7; 1.04: 3.3574; 1.05: 3.548; 1.06: 6; 1.07: 3.067; 1.08: 3.67 |  | 1
| 1 | 2 | 1.01: 3.8; 1.02: 3.33; 1.03: 3.5; 1.04: 3.3093; 1.05: 3.552; 1.06: 4; 1.07: 3.0428; 1.08: 3.62 |  | 7
| 1 | 3 | 1.01: 3.5; 1.02: 3.43; 1.03: 3.7; 1.04: 3.4301; 1.05: 3.546; 1.06: 4; 1.07: 2.9943; 1.08: 3.77 |  | 16
| 1 | 4 | 1.01: 3.1; 1.02: 3.59; 1.03: 3.5; 1.04: 3.1534; 1.05: 3.514; 1.06: 4; 1.07: 3.074; 1.08: 3.64 |  | 3.4207
| 1 | 5 | 1.01: 3.4; 1.02: 3.32; 1.03: 3.7; 1.04: 3.3741; 1.05: 3.537; 1.06: 5; 1.07: 3.1944; 1.08: 3.89 |  | 130
| 1 | 6 | 1.01: 3; 1.02: 3.53; 1.03: 3.2; 1.04: 3.4936; 1.05: 3.548; 1.06: 4; 1.07: 2.9659; 1.08: 3.78
| Mean |  | 1.01: 3.3167; 1.02: 3.4; 1.03: 3.55; 1.04: 3.353; 1.05: 3.5408; 1.06: 4.5; 1.07: 3.0564; 1.08: 3.7283
| Median |  | 1.01: 3.25; 1.02: 3.38; 1.03: 3.6; 1.04: 3.3658; 1.05: 3.547; 1.06: 4; 1.07: 3.0549; 1.08: 3.72
| Std Dev. |  | 1.01: 0.30605; 1.02: 0.14505; 1.03: 0.19748; 1.04: 0.11656; 1.05: 0.014063; 1.06: 0.83666; 1.07: 0.079609; 1.08: 0.10342
| Rel.Std.Dev. |  | 1.01: 9.228%; 1.02: 4.266%; 1.03: 5.563%; 1.04: 3.476%; 1.05: 0.3972%; 1.06: 18.59%; 1.07: 2.605%; 1.08: 2.774%
| PDM3 |  | 1.01: -3.043%; 1.02: -0.6065%; 1.03: 3.779%; 1.04: -1.981%; 1.05: 3.511%; 1.06: 31.55%; 1.07: -10.65%; 1.08: 8.992%
| Z-Score (Absolute) |  | 1.01: 0.4; 1.02: 0; 1.03: 0.72; 1.04: 0.23; 1.05: 0.67; 1.06: NA; 1.07: 1.65; 1.08: 1.57
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Er in OREAS 627 (Certified Value 1.69 ppm) | [24 empty] | Certified
Er | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 2
3
| 1 | 1 | 1.01: 1.8; 1.02: 1.52; 1.03: 1.7; 1.04: 1.7471; 1.05: 1.767; 1.06: 3; 1.07: 1.7673; 1.08: 1.65 |  | 1
| 1 | 2 | 1.01: 2; 1.02: 1.61; 1.03: 1.6; 1.04: 1.5577; 1.05: 1.753; 1.06: < 2; 1.07: 1.5267; 1.08: 1.66 |  | 8
| 1 | 3 | 1.01: 1.9; 1.02: 1.65; 1.03: 1.7; 1.04: 1.6685; 1.05: 1.742; 1.06: < 2; 1.07: 1.4108; 1.08: 1.75 |  | 16
| 1 | 4 | 1.01: 1.2; 1.02: 1.68; 1.03: 1.7; 1.04: 1.6534; 1.05: 1.754; 1.06: 2; 1.07: 1.6347; 1.08: 1.76 |  | 1.6874
| 1 | 5 | 1.01: 1.8; 1.02: 1.54; 1.03: 1.6; 1.04: 1.8764; 1.05: 1.756; 1.06: 2; 1.07: 1.5783; 1.08: 1.76 |  | 131
| 1 | 6 | 1.01: 1.7; 1.02: 1.63; 1.03: 1.4; 1.04: 1.695; 1.05: 1.762; 1.06: < 2; 1.07: 1.5712; 1.08: 1.7
| Mean |  | 1.01: 1.7333; 1.02: 1.605; 1.03: 1.6167; 1.04: 1.6997; 1.05: 1.7557; 1.06: 2.3333; 1.07: 1.5815; 1.08: 1.7133
| Median |  | 1.01: 1.8; 1.02: 1.62; 1.03: 1.65; 1.04: 1.6818; 1.05: 1.755; 1.06: 2; 1.07: 1.5747; 1.08: 1.725
| Std Dev. |  | 1.01: 0.28048; 1.02: 0.062849; 1.03: 0.1169; 1.04: 0.10654; 1.05: 0.0085479; 1.06: 0.57735; 1.07: 0.11799; 1.08: 0.050465
| Rel.Std.Dev. |  | 1.01: 16.18%; 1.02: 3.916%; 1.03: 7.231%; 1.04: 6.268%; 1.05: 0.4869%; 1.06: 24.74%; 1.07: 7.461%; 1.08: 2.945%
| PDM3 |  | 1.01: 2.722%; 1.02: -4.884%; 1.03: -4.192%; 1.04: 0.7279%; 1.05: 4.045%; 1.06: 38.28%; 1.07: -6.277%; 1.08: 1.537%
| Z-Score (Absolute) |  | 1.01: 0.58; 1.02: 0.9; 1.03: 0.77; 1.04: 0.19; 1.05: 0.84; 1.06: 0.19; 1.07: 1.18; 1.08: 0.35
| Analytical results for Eu in OREAS 627 (Certified Value 1.42 ppm) | [24 empty] | Certified
Eu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS |  | 2
2
| 1 | 1 | 1.01: 2.1; 1.02: 1.21; 1.03: 1.3; 1.04: 1.8317; 1.05: 1.53; 1.06: < 2; 1.07: 1.27 |  | 1
| 1 | 2 | 1.01: 2.1; 1.02: 1.28; 1.03: 1.4; 1.04: 1.684; 1.05: 1.499; 1.06: < 2; 1.07: 1.22 |  | 9
| 1 | 3 | 1.01: 2.4; 1.02: 1.28; 1.03: 1.4; 1.04: 1.7752; 1.05: 1.486; 1.06: < 2; 1.07: 1.31 |  | 16
| 1 | 4 | 1.01: 2.1; 1.02: 1.3; 1.03: 1.4; 1.04: 1.7362; 1.05: 1.523; 1.06: < 2; 1.07: 1.28 |  | 1.4216
| 1 | 5 | 1.01: 2.3; 1.02: 1.2; 1.03: 1.1; 1.04: 1.8941; 1.05: 1.531; 1.06: < 2; 1.07: 1.28 |  | 132
| 1 | 6 | 1.01: 2.1; 1.02: 1.32; 1.03: 1.1; 1.04: 1.7442; 1.05: 1.535; 1.06: < 2; 1.07: 1.23
| Mean |  | 1.01: 2.1833; 1.02: 1.265; 1.03: 1.2833; 1.04: 1.7776; 1.05: 1.5173; 1.07: 1.265
| Median |  | 1.01: 2.1; 1.02: 1.28; 1.03: 1.35; 1.04: 1.7597; 1.05: 1.5265; 1.07: 1.275
| Std Dev. |  | 1.01: 0.13292; 1.02: 0.048888; 1.03: 0.1472; 1.04: 0.074959; 1.05: 0.020047; 1.07: 0.033912
| Rel.Std.Dev. |  | 1.01: 6.088%; 1.02: 3.865%; 1.03: 11.47%; 1.04: 4.217%; 1.05: 1.321%; 1.07: 2.681%
| PDM3 |  | 1.01: 53.58%; 1.02: -11.02%; 1.03: -9.729%; 1.04: 25.04%; 1.05: 6.731%; 1.07: -11.02%
| Z-Score (Absolute) |  | 1.01: 2.59; 1.02: 0.05; 1.03: 0; 1.04: 1.42; 1.05: 0.67; 1.06: 0.82; 1.07: 0.05
| Analytical results for Fe in OREAS 627 (Certified Value 6.82 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 6.95 | 6.96 | 6.2 | 6.638 | 6.617 | 6.841 | 6.722 | 6.589 | 7.39 | 6.89 | 6.73 | 6.86 | 6.2971 | 7.4095 | 6.87 | 6.88 | 7.4153 | 7.3624 | 6.75 | 7.47 | 6.44 | 6.69 | 7.21 | 1
| 1 | 2 | 7.01 | 7.06 | 6.28 | 6.701 | 6.729 | 6.89 | 6.729 | 6.736 | 7.36 | 6.79 | 6.79 | 6.75 | 6.2565 | 7.4623 | 6.99 | 6.83 | 7.4881 | 7.3205 | 6.79 | 7.49 | 6.49 | 6.88 | 7.23 | 26
| 1 | 3 | 6.96 | 7.02 | 6.14 | 6.729 | 6.729 | 6.848 | 6.708 | 6.827 | 7.46 | 6.82 | 6.72 | 6.8 | 6.2402 | 7.3213 | 6.96 | 6.88 | 7.3596 | 7.3354 | 6.78 | 7.45 | 6.39 | 6.81 | 7.22 | 16
| 1 | 4 | 6.93 | 6.97 | 6.12 | 6.715 | 6.61 | 6.939 | 6.722 | 6.624 | 7.44 | 6.9 | 7.04 | 6.61 | 6.3021 | 7.4425 | 6.87 | 6.83 | 7.5795 | 7.3668 | 6.73 | 7.43 | 6.44 | 6.89 | 7.24 | 6.8179
| 1 | 5 | 7.01 | 7.01 | 6.29 | 6.708 | 6.771 | 6.939 | 6.701 | 6.638 | 7.49 | 6.68 | 6.69 | 6.92 | 6.3765 | 7.294 | 6.99 | 6.88 | 7.6219 | 7.3686 | 6.78 | 7.4 | 6.51 | 6.81 | 7.17 | 133
| 1 | 6 | 6.98 | 7.22 | 6.17 | 6.68 | 6.827 | 6.848 | 6.883 | 6.673 | 7.39 | 6.8 | 6.72 | 6.72 | 6.3356 | 7.2864 | 6.93 | 6.83 | 7.2592 | 7.3867 | 6.8 | 7.4 | 6.45 | 6.74 | 7.24
| Mean |  | 6.9733 | 7.04 | 6.2 | 6.6952 | 6.7138 | 6.8842 | 6.7442 | 6.6812 | 7.4217 | 6.8133 | 6.7817 | 6.7767 | 6.3013 | 7.3693 | 6.935 | 6.855 | 7.4539 | 7.3567 | 6.7717 | 7.44 | 6.4533 | 6.8033 | 7.2183
| Median |  | 6.97 | 7.015 | 6.185 | 6.7045 | 6.729 | 6.869 | 6.722 | 6.6555 | 7.415 | 6.81 | 6.725 | 6.775 | 6.2996 | 7.3654 | 6.945 | 6.855 | 7.4517 | 7.3646 | 6.78 | 7.44 | 6.445 | 6.81 | 7.225
| Std Dev. |  | 0.03266 | 0.095289 | 0.071274 | 0.032357 | 0.085665 | 0.045884 | 0.068788 | 0.08714 | 0.049565 | 0.079917 | 0.13075 | 0.1093 | 0.050181 | 0.078082 | 0.055045 | 0.027386 | 0.13672 | 0.024273 | 0.026394 | 0.036878 | 0.042269 | 0.077889 | 0.026394
| Rel.Std.Dev. |  | 0.4684% | 1.354% | 1.15% | 0.4833% | 1.276% | 0.6665% | 1.02% | 1.304% | 0.6678% | 1.173% | 1.928% | 1.613% | 0.7964% | 1.06% | 0.7937% | 0.3995% | 1.834% | 0.3299% | 0.3898% | 0.4957% | 0.655% | 1.145% | 0.3657%
| PDM3 |  | 2.279% | 3.257% | -9.063% | -1.801% | -1.527% | 0.9715% | -1.082% | -2.006% | 8.855% | -0.06741% | -0.5319% | -0.6052% | -7.577% | 8.088% | 1.717% | 0.5437% | 9.328% | 7.903% | -0.6785% | 9.124% | -5.348% | -0.2141% | 5.873%
| Z-Score (Absolute) |  | 0.81 | 1.15 | 3.12 | 0.6 | 0.51 | 0.36 | 0.35 | 0.67 | 3.1 | 0 | 0.16 | 0.19 | 2.61 | 2.83 | 0.62 | 0.21 | 3.26 | 2.77 | 0.21 | 3.19 | 1.83 | 0.05 | 2.06
| Analytical results for Ga in OREAS 627 (Certified Value 18.3 ppm) | [24 empty] | Certified
Ga | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.16; 1.08: 1.17; 1.09: 1.20; 1.10: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*OES/MS; 1.09: PF*MS; 1.10: PF*MS | 1
2
| 1 | 1 | 1.01: 18; 1.02: 19.7; 1.03: 16.4; 1.04: 17; 1.05: 19.7306; 1.06: 20.6615; 1.07: 132.9; 1.08: 19.7; 1.09: 16; 1.10: 19 | 1
| 1 | 2 | 1.01: 18; 1.02: 17.9; 1.03: 17.1; 1.04: 18; 1.05: 18.96; 1.06: 20.1142; 1.07: 132.3; 1.08: 20.2; 1.09: 16; 1.10: 19 | 44
| 1 | 3 | 1.01: 18; 1.02: 18.9; 1.03: 17.6; 1.04: 18; 1.05: 19.242; 1.06: 20.2518; 1.07: 132.9; 1.08: 20.4; 1.09: 17; 1.10: 19 | 16
| 1 | 4 | 1.01: 17; 1.02: 16.6; 1.03: 18.4; 1.04: 16; 1.05: 19.1944; 1.06: 20.4469; 1.07: 132.9; 1.08: 19.7; 1.09: 16; 1.10: 19 | 18.347
| 1 | 5 | 1.01: 18; 1.02: 17.4; 1.03: 16.4; 1.04: 14; 1.05: 19.1851; 1.06: 20.1157; 1.07: 131.4; 1.08: 20.9; 1.09: 17; 1.10: 20 | 134
| 1 | 6 | 1.01: 18; 1.02: 18; 1.03: 18.2; 1.04: 15; 1.05: 19.3313; 1.06: 20.5056; 1.07: 129.6; 1.08: 20.5; 1.09: 16; 1.10: 20
| Mean |  | 1.01: 17.8333; 1.02: 18.0833; 1.03: 17.35; 1.04: 16.3333; 1.05: 19.2739; 1.06: 20.3493; 1.07: 132; 1.08: 20.2333; 1.09: 16.3333; 1.10: 19.3333
| Median |  | 1.01: 18; 1.02: 17.95; 1.03: 17.35; 1.04: 16.5; 1.05: 19.2182; 1.06: 20.3493; 1.07: 132.6; 1.08: 20.3; 1.09: 16; 1.10: 19
| Std Dev. |  | 1.01: 0.40825; 1.02: 1.0944; 1.03: 0.8666; 1.04: 1.633; 1.05: 0.25519; 1.06: 0.22393; 1.07: 1.3145; 1.08: 0.47188; 1.09: 0.5164; 1.10: 0.5164
| Rel.Std.Dev. |  | 1.01: 2.289%; 1.02: 6.052%; 1.03: 4.995%; 1.04: 9.998%; 1.05: 1.324%; 1.06: 1.1%; 1.07: 0.9959%; 1.08: 2.332%; 1.09: 3.162%; 1.10: 2.671%
| PDM3 |  | 1.01: -2.8%; 1.02: -1.437%; 1.03: -5.434%; 1.04: -10.98%; 1.05: 5.052%; 1.06: 10.91%; 1.07: 619.5%; 1.08: 10.28%; 1.09: -10.98%; 1.10: 5.376%
| Z-Score (Absolute) |  | 1.01: 0.4; 1.02: 0.28; 1.03: 0.62; 1.04: 1.1; 1.05: 0.28; 1.06: 0.78; 1.07: 53; 1.08: 0.73; 1.09: 1.1; 1.10: 0.31
| Analytical results for Gd in OREAS 627 (Certified Value 4.99 ppm) | [24 empty] | Certified
Gd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | 2
3
| 1 | 1 | 1.01: 5.9; 1.02: 4.57; 1.03: 4.7; 1.04: 4.9039; 1.05: 4.943; 1.06: 5; 1.07: 5.2783; 1.08: 4.9 | 1
| 1 | 2 | 1.01: 5.6; 1.02: 5; 1.03: 4.6; 1.04: 4.7174; 1.05: 4.968; 1.06: 4; 1.07: 5.1975; 1.08: 4.87 | 10
| 1 | 3 | 1.01: 4.8; 1.02: 4.78; 1.03: 4.9; 1.04: 4.7401; 1.05: 5.017; 1.06: 4; 1.07: 4.9996; 1.08: 5.14 | 16
| 1 | 4 | 1.01: 4.9; 1.02: 5.08; 1.03: 4.7; 1.04: 5.0362; 1.05: 4.997; 1.06: 4; 1.07: 5.1868; 1.08: 4.76 | 4.9917
| 1 | 5 | 1.01: 5.9; 1.02: 4.68; 1.03: 4.4; 1.04: 5.0711; 1.05: 4.963; 1.06: 4; 1.07: 5.4398; 1.08: 5.18 | 135
| 1 | 6 | 1.01: 4.9; 1.02: 4.96; 1.03: 4.7; 1.04: 4.9171; 1.05: 4.997; 1.06: 4; 1.07: 5.3279; 1.08: 5.03
| Mean |  | 1.01: 5.3333; 1.02: 4.845; 1.03: 4.6667; 1.04: 4.8976; 1.05: 4.9808; 1.06: 4.1667; 1.07: 5.2383; 1.08: 4.98
| Median |  | 1.01: 5.25; 1.02: 4.87; 1.03: 4.7; 1.04: 4.9105; 1.05: 4.9825; 1.06: 4; 1.07: 5.2379; 1.08: 4.965
| Std Dev. |  | 1.01: 0.52409; 1.02: 0.19977; 1.03: 0.1633; 1.04: 0.14629; 1.05: 0.027367; 1.06: 0.40825; 1.07: 0.14937; 1.08: 0.16432
| Rel.Std.Dev. |  | 1.01: 9.827%; 1.02: 4.123%; 1.03: 3.499%; 1.04: 2.987%; 1.05: 0.5495%; 1.06: 9.798%; 1.07: 2.851%; 1.08: 3.3%
| PDM3 |  | 1.01: 6.844%; 1.02: -2.939%; 1.03: -6.511%; 1.04: -1.884%; 1.05: -0.2173%; 1.06: -16.53%; 1.07: 4.941%; 1.08: -0.234%
| Z-Score (Absolute) |  | 1.01: 1.76; 1.02: 0.67; 1.03: 1.57; 1.04: 0.41; 1.05: 0; 1.06: NA; 1.07: 1.29; 1.08: 0
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Ge in OREAS 627 (Certified Value 2.22 ppm) | [24 empty] | Certified
Ge | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.17; 1.06: 1.20; 1.07: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS | 2
2
| 1 | 1 | 1.01: 2.1; 1.02: 2; 1.03: 2; 1.04: 2.4134; 1.05: 3.8; 1.06: < 2; 1.07: 2 | 1
| 1 | 2 | 1.01: 2.3; 1.02: 3; 1.03: 2; 1.04: 2.3054; 1.05: 3.5; 1.06: < 2; 1.07: 2 | 46
| 1 | 3 | 1.01: 2.1; 1.02: 3; 1.03: 2; 1.04: 2.2918; 1.05: 3.6; 1.06: < 2; 1.07: 2 | 16
| 1 | 4 | 1.01: 2.6; 1.02: 3; 1.03: 2; 1.04: 2.2242; 1.05: 3.4; 1.06: < 2; 1.07: 2 | 2.2228
| 1 | 5 | 1.01: 2.1; 1.02: 3; 1.03: 2; 1.04: 2.4174; 1.05: 3.7; 1.06: < 2; 1.07: 2 | 136
| 1 | 6 | 1.01: 1.9; 1.02: 2; 1.03: 2; 1.04: 2.4323; 1.05: 3.5; 1.06: < 2; 1.07: 2
| Mean |  | 1.01: 2.1833; 1.02: 2.6667; 1.03: 2; 1.04: 2.3474; 1.05: 3.5833; 1.07: 2
| Median |  | 1.01: 2.1; 1.02: 3; 1.03: 2; 1.04: 2.3594; 1.05: 3.55; 1.07: 2
| Std Dev. |  | 1.01: 0.24014; 1.02: 0.5164; 1.03: 0; 1.04: 0.085444; 1.05: 0.1472; 1.07: 0
| Rel.Std.Dev. |  | 1.01: 11%; 1.02: 19.36%; 1.03: 0%; 1.04: 3.64%; 1.05: 4.108%; 1.07: 0%
| PDM3 |  | 1.01: -1.776%; 1.02: 19.97%; 1.03: -10.02%; 1.04: 5.606%; 1.05: 61.21%; 1.07: -10.02%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 1.78; 1.03: 0.67; 1.04: 0.6; 1.05: 5.15; 1.06: 4.35; 1.07: 0.67
| Analytical results for Hf in OREAS 627 (Indicative Value 4.46 ppm) | [24 empty] | Indicative
Hf | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.10; 1.03: 1.13; 1.04: 1.20; 1.05: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*MS | 2
2
| 1 | 1 | 1.01: 20; 1.02: 4.6; 1.03: 5.4876; 1.04: 4; 1.05: 4 | 1
| 1 | 2 | 1.01: < 10; 1.02: 4.6; 1.03: 5.0297; 1.04: 4; 1.05: 4 | 16
| 1 | 3 | 1.01: < 10; 1.02: 5; 1.03: 5.3374; 1.04: 4; 1.05: 4 | 16
| 1 | 4 | 1.01: < 10; 1.02: 4.7; 1.03: 5.0752; 1.04: 4; 1.05: 4 | 4.456
| 1 | 5 | 1.01: < 10; 1.02: 4.8; 1.03: 5.2992; 1.04: 4; 1.05: 4 | 22
| 1 | 6 | 1.01: < 10; 1.02: 4.2; 1.03: 4.8145; 1.04: 4; 1.05: 4
| Mean |  | 1.01: 20; 1.02: 4.65; 1.03: 5.1739; 1.04: 4; 1.05: 4
| Median |  | 1.01: 20; 1.02: 4.65; 1.03: 5.1872; 1.04: 4; 1.05: 4
| Std Dev. |  | 1.02: 0.26646; 1.03: 0.24517; 1.04: 0; 1.05: 0
| Rel.Std.Dev. |  | 1.02: 5.73%; 1.03: 4.739%; 1.04: 0%; 1.05: 0%
| PDM3 |  | 1.01: 348.8%; 1.02: 4.354%; 1.03: 16.11%; 1.04: -10.23%; 1.05: -10.23%
| Z-Score (Absolute) |  | 1.01: 2.96; 1.02: 0; 1.03: 0.54; 1.04: 0.67; 1.05: 0.67
| Analytical results for Hg in OREAS 627 (Indicative Value < 5 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | 1.01: Lab | 1
| No | No | 1.01: 1.16 | ppm
|  |  | 1.01: PF*OES | 2
2
| 1 | 1 | 1.01: < 5 | 1
| 1 | 2 | 1.01: < 5 | 4
| 1 | 3 | 1.01: < 5 | 16
| 1 | 4 | 1.01: < 5 | < 5
| 1 | 5 | 1.01: < 5 | 18
| 1 | 6 | 1.01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01
| Analytical results for Ho in OREAS 627 (Certified Value 0.63 ppm) | [24 empty] | Certified
Ho | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | 2
3
| 1 | 1 | 1.01: 0.7; 1.02: 0.57; 1.03: 0.6; 1.04: 0.62658; 1.05: 0.676; 1.06: < 2; 1.07: 0.63065; 1.08: 0.64 | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.6; 1.03: 0.7; 1.04: 0.58139; 1.05: 0.681; 1.06: < 2; 1.07: 0.53849; 1.08: 0.62 | 6
| 1 | 3 | 1.01: 0.8; 1.02: 0.61; 1.03: 0.6; 1.04: 0.6218; 1.05: 0.684; 1.06: < 2; 1.07: 0.48723; 1.08: 0.66 | 16
| 1 | 4 | 1.01: 0.5; 1.02: 0.63; 1.03: 0.6; 1.04: 0.60587; 1.05: 0.682; 1.06: < 2; 1.07: 0.63318; 1.08: 0.66 | 0.63142
| 1 | 5 | 1.01: 0.7; 1.02: 0.57; 1.03: 0.7; 1.04: 0.66926; 1.05: 0.678; 1.06: < 2; 1.07: 0.56053; 1.08: 0.67 | 137
| 1 | 6 | 1.01: 0.6; 1.02: 0.63; 1.03: 0.6; 1.04: 0.61882; 1.05: 0.679; 1.06: < 2; 1.07: 0.52601; 1.08: 0.68
| Mean |  | 1.01: 0.66667; 1.02: 0.60167; 1.03: 0.63333; 1.04: 0.62062; 1.05: 0.68; 1.07: 0.56268; 1.08: 0.655
| Median |  | 1.01: 0.7; 1.02: 0.605; 1.03: 0.6; 1.04: 0.62031; 1.05: 0.68; 1.07: 0.54951; 1.08: 0.66
| Std Dev. |  | 1.01: 0.10328; 1.02: 0.027142; 1.03: 0.05164; 1.04: 0.028852; 1.05: 0.0028983; 1.07: 0.058681; 1.08: 0.021679
| Rel.Std.Dev. |  | 1.01: 15.49%; 1.02: 4.511%; 1.03: 8.154%; 1.04: 4.649%; 1.05: 0.4262%; 1.07: 10.43%; 1.08: 3.31%
| PDM3 |  | 1.01: 5.581%; 1.02: -4.713%; 1.03: 0.3024%; 1.04: -1.711%; 1.05: 7.693%; 1.07: -10.89%; 1.08: 3.734%
| Z-Score (Absolute) |  | 1.01: 0.51; 1.02: 0.97; 1.03: 0.25; 1.04: 0.53; 1.05: 0.81; 1.06: 8.08; 1.07: 1.85; 1.08: 0.25
| Analytical results for In in OREAS 627 (Certified Value 9.12 ppm) | [24 empty] | Certified
In | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.17; 1.06: 1.20; 1.07: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS | 2
3
| 1 | 1 | 1.01: 8.3; 1.02: 8; 1.03: 9.1; 1.04: 9.5343; 1.05: 9.57; 1.06: 10.3; 1.07: 8.5 | 1
| 1 | 2 | 1.01: 8.4; 1.02: 8.3; 1.03: 9.4; 1.04: 9.1113; 1.05: 9.72; 1.06: 10.3; 1.07: 8.6 | 33
| 1 | 3 | 1.01: 8.3; 1.02: 8.7; 1.03: 9.6; 1.04: 9.1955; 1.05: 9.71; 1.06: 10.1; 1.07: 8.8 | 16
| 1 | 4 | 1.01: 8.6; 1.02: 8.7; 1.03: 9.6; 1.04: 9.4335; 1.05: 9.22; 1.06: 10.1; 1.07: 8.7 | 9.1184
| 1 | 5 | 1.01: 9; 1.02: 8; 1.03: 9.2; 1.04: 8.9385; 1.05: 9.86; 1.06: 10.2; 1.07: 9 | 138
| 1 | 6 | 1.01: 8.3; 1.02: 8.7; 1.03: 9.1; 1.04: 8.7901; 1.05: 9.41; 1.06: 10.2; 1.07: 9
| Mean |  | 1.01: 8.4833; 1.02: 8.4; 1.03: 9.3333; 1.04: 9.1672; 1.05: 9.5817; 1.06: 10.2; 1.07: 8.7667
| Median |  | 1.01: 8.35; 1.02: 8.5; 1.03: 9.3; 1.04: 9.1534; 1.05: 9.64; 1.06: 10.2; 1.07: 8.75
| Std Dev. |  | 1.01: 0.27869; 1.02: 0.34641; 1.03: 0.23381; 1.04: 0.28428; 1.05: 0.23387; 1.06: 0.089443; 1.07: 0.20656
| Rel.Std.Dev. |  | 1.01: 3.285%; 1.02: 4.124%; 1.03: 2.505%; 1.04: 3.101%; 1.05: 2.441%; 1.06: 0.8769%; 1.07: 2.356%
| PDM3 |  | 1.01: -6.965%; 1.02: -7.879%; 1.03: 2.357%; 1.04: 0.5353%; 1.05: 5.08%; 1.06: 11.86%; 1.07: -3.858%
| Z-Score (Absolute) |  | 1.01: 1.11; 1.02: 1.25; 1.03: 0.27; 1.04: 0; 1.05: 0.67; 1.06: 1.68; 1.07: 0.65
| Analytical results for K in OREAS 627 (Certified Value 2.56 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 2
3
| 1 | 1 | 2.74 | 2.6 | 0.4 | 2.482 | 2.549 | 2.532 | 2.54 | 2.648 | 2.57 | 2.6 | 2.57 | 2.38 | 2.5487 | 2.756 | 2.55 | 2.76 | 2.5552 | 2.5 | 2.5 | 2.56 | 2.5 |  |  | 1
| 1 | 2 | 2.75 | 2.5 | 0.4 | 2.499 | 2.557 | 2.59 | 2.524 | 2.64 | 2.56 | 2.53 | 2.65 | 2.36 | 2.5512 | 2.7797 | 2.61 | 2.76 | 2.5821 | 2.6 | 2.5 | 2.59 | 2.6 |  |  | 0
| 1 | 3 | 2.79 | 2.5 | 0.4 | 2.507 | 2.59 | 2.623 | 2.549 | 2.656 | 2.58 | 2.59 | 2.72 | 2.41 | 2.4847 | 2.8264 | 2.59 | 2.78 | 2.5222 | 2.5 | 2.5 | 2.59 | 2.5 |  |  | 16
| 1 | 4 | 2.74 | 2.5 | 0.4 | 2.474 | 2.573 | 2.607 | 2.549 | 2.573 | 2.57 | 2.61 | 2.75 | 2.33 | 2.5148 | 2.7869 | 2.55 | 2.78 | 2.641 | 2.5 | 2.6 | 2.56 | 2.6 |  |  | 2.5635
| 1 | 5 | 2.78 | 2.5 | 0.5 | 2.482 | 2.507 | 2.573 | 2.54 | 2.632 | 2.59 | 2.57 | 2.66 | 2.44 | 2.5329 | 2.8247 | 2.59 | 2.76 | 2.6694 | 2.6 | 2.6 | 2.57 | 2.5 |  |  | 139
| 1 | 6 | 2.74 | 2.6 | 0.4 | 2.49 | 2.623 | 2.557 | 2.549 | 2.64 | 2.56 | 2.57 | 2.59 | 2.39 | 2.4981 | 2.7711 | 2.6 | 2.75 | 2.4789 | 2.6 | 2.5 | 2.57 | 2.6
| Mean |  | 2.7567 | 2.5333 | 0.41667 | 2.489 | 2.5665 | 2.5803 | 2.5418 | 2.6315 | 2.5717 | 2.5783 | 2.6567 | 2.385 | 2.5217 | 2.7908 | 2.5817 | 2.765 | 2.5748 | 2.55 | 2.5333 | 2.5733 | 2.55
| Median |  | 2.745 | 2.5 | 0.4 | 2.486 | 2.565 | 2.5815 | 2.5445 | 2.64 | 2.57 | 2.58 | 2.655 | 2.385 | 2.5239 | 2.7833 | 2.59 | 2.76 | 2.5686 | 2.55 | 2.5 | 2.57 | 2.55
| Std Dev. |  | 0.022509 | 0.05164 | 0.040825 | 0.012231 | 0.039303 | 0.033345 | 0.009786 | 0.029798 | 0.01169 | 0.028577 | 0.070333 | 0.038341 | 0.027195 | 0.028835 | 0.025626 | 0.012247 | 0.071739 | 0.054772 | 0.05164 | 0.013663 | 0.054772
| Rel.Std.Dev. |  | 0.8165% | 2.038% | 9.798% | 0.4914% | 1.531% | 1.292% | 0.385% | 1.132% | 0.4546% | 1.108% | 2.647% | 1.608% | 1.078% | 1.033% | 0.9926% | 0.4429% | 2.786% | 2.148% | 2.038% | 0.5309% | 2.148%
| PDM3 |  | 7.537% | -1.175% | -83.75% | -2.905% | 0.1184% | 0.6581% | -0.8438% | 2.654% | 0.32% | 0.58% | 3.636% | -6.962% | -1.628% | 8.868% | 0.7101% | 7.862% | 0.4422% | -0.5252% | -1.175% | 0.385% | -0.5252%
| Z-Score (Absolute) |  | 3.25 | 0.67 | 37.91 | 1.45 | 0.09 | 0.15 | 0.53 | 1.05 | 0 | 0.12 | 1.5 | 3.28 | 0.88 | 3.85 | 0.18 | 3.4 | 0.06 | 0.38 | 0.67 | 0.03 | 0.38
| Analytical results for La in OREAS 627 (Certified Value 34.5 ppm) | [24 empty] | Certified
La | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.16; 1.07: 1.19; 1.08: 1.20; 1.09: 1.23; 1.10: 1.25; 1.11: 1.26 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*OES; 1.07: BF*MS; 1.08: PF*MS; 1.09: PF*MS; 1.10: PF*OES; 1.11: PF*OES |  |  | 1
2
| 1 | 1 | 1.01: 35; 1.02: 35.2; 1.03: 31.6; 1.04: 36.4; 1.05: 35.6938; 1.06: 36.9; 1.07: 33.61; 1.08: 53; 1.09: 32.9; 1.10: 33; 1.11: 30 |  |  | 1
| 1 | 2 | 1.01: 36; 1.02: 36.4; 1.03: 33.2; 1.04: 36.2; 1.05: 34.1314; 1.06: 34.7; 1.07: 33.634; 1.08: 33; 1.09: 31.9; 1.10: 32; 1.11: 30 |  |  | 34
| 1 | 3 | 1.01: 36; 1.02: 36.4; 1.03: 33.3; 1.04: 36.9; 1.05: 34.2918; 1.06: 35.4; 1.07: 33.693; 1.08: 32; 1.09: 33.4; 1.10: 34; 1.11: 30 |  |  | 16
| 1 | 4 | 1.01: 36; 1.02: 33.7; 1.03: 34.9; 1.04: 37; 1.05: 34.896; 1.06: 35.2; 1.07: 33.644; 1.08: 35; 1.09: 33.9; 1.10: 33; 1.11: 30 |  |  | 34.4574
| 1 | 5 | 1.01: 36; 1.02: 36.8; 1.03: 31.9; 1.04: 36.6; 1.05: 36.287; 1.06: 35.1; 1.07: 33.657; 1.08: 42; 1.09: 34.7; 1.10: 34; 1.11: 30 |  |  | 140
| 1 | 6 | 1.01: 36; 1.02: 35.5; 1.03: 34; 1.04: 35.9; 1.05: 35.3974; 1.06: 35.3; 1.07: 33.668; 1.08: 32; 1.09: 33.3; 1.10: 33; 1.11: 30
| Mean |  | 1.01: 35.8333; 1.02: 35.6667; 1.03: 33.15; 1.04: 36.5; 1.05: 35.1162; 1.06: 35.4333; 1.07: 33.651; 1.08: 37.8333; 1.09: 33.35; 1.10: 33.1667; 1.11: 30
| Median |  | 1.01: 36; 1.02: 35.95; 1.03: 33.25; 1.04: 36.5; 1.05: 35.1467; 1.06: 35.25; 1.07: 33.6505; 1.08: 34; 1.09: 33.35; 1.10: 33; 1.11: 30
| Std Dev. |  | 1.01: 0.40825; 1.02: 1.1378; 1.03: 1.247; 1.04: 0.41952; 1.05: 0.83443; 1.06: 0.75807; 1.07: 0.028664; 1.08: 8.3287; 1.09: 0.94181; 1.10: 0.75277; 1.11: 0
| Rel.Std.Dev. |  | 1.01: 1.139%; 1.02: 3.19%; 1.03: 3.762%; 1.04: 1.149%; 1.05: 2.376%; 1.06: 2.139%; 1.07: 0.08518%; 1.08: 22.01%; 1.09: 2.824%; 1.10: 2.27%; 1.11: 0%
| PDM3 |  | 1.01: 3.993%; 1.02: 3.509%; 1.03: -3.794%; 1.04: 5.928%; 1.05: 1.912%; 1.06: 2.832%; 1.07: -2.34%; 1.08: 9.797%; 1.09: -3.214%; 1.10: -3.746%; 1.11: -12.94%
| Z-Score (Absolute) |  | 1.01: 0.26; 1.02: 0.19; 1.03: 1.01; 1.04: 0.58; 1.05: 0.08; 1.06: 0.08; 1.07: 0.77; 1.08: 1.21; 1.09: 0.91; 1.10: 1; 1.11: NA
Results from laboratory 1.26 were removed due to their 10 ppm reading resolution.
| Analytical results for Li in OREAS 627 (Certified Value 20.1 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 1
1
| 1 | 1 | 17 | 22 | < 10 | 19 | 20 | 20 | < 100 | 23.4996 | 19.2 | 15 | 26 | 22 | 20 | < 100 | [9 empty] | 1
| 1 | 2 | 17 | 16 | < 10 | 20 | 19 | 20 | < 100 | 33.7916 | 18.1 | 15 | 22 | 26 | 30 | < 100 | [9 empty] | 0
| 1 | 3 | 18 | 21 | < 10 | 20 | 19 | 20 | < 100 | 19.502 | 18.9 | 15 | 18 | 25 | 30 | < 100 | [9 empty] | 16
| 1 | 4 | 18 | 17 | < 10 | 19 | 19 | 20 | < 100 | 24.3355 | 18.4 | 14 | 27 | 25 | 30 | < 100 | [9 empty] | 20.0726
| 1 | 5 | 17 | 19 | < 10 | 20 | 19 | 20 | < 100 | 28.4982 | 19 | 15 | 24 | 22 | 30 | < 100 | [9 empty] | 141
| 1 | 6 | 18 | 21 | < 10 | 20 | 19 | 20 | < 100 | 36.1376 | 18 | 14 | 24 | 26 | 30 | < 100
| Mean |  | 1.01: 17.5; 1.02: 19.3333; 1.04: 19.6667; 1.05: 19.1667; 1.06: 20; 1.08: 27.6274; 1.09: 18.6; 1.10: 14.6667; 1.11: 23.5; 1.12: 24.3333; 1.13: 28.3333
| Median |  | 1.01: 17.5; 1.02: 20; 1.04: 20; 1.05: 19; 1.06: 20; 1.08: 26.4168; 1.09: 18.65; 1.10: 15; 1.11: 24; 1.12: 25; 1.13: 30
| Std Dev. |  | 1.01: 0.54772; 1.02: 2.4221; 1.04: 0.5164; 1.05: 0.40825; 1.06: 0; 1.08: 6.4043; 1.09: 0.502; 1.10: 0.5164; 1.11: 3.2094; 1.12: 1.8619; 1.13: 4.0825
| Rel.Std.Dev. |  | 1.01: 3.13%; 1.02: 12.53%; 1.04: 2.626%; 1.05: 2.13%; 1.06: 0%; 1.08: 23.18%; 1.09: 2.699%; 1.10: 3.521%; 1.11: 13.66%; 1.12: 7.652%; 1.13: 14.41%
| PDM3 |  | 1.01: -12.82%; 1.02: -3.683%; 1.04: -2.022%; 1.05: -4.513%; 1.06: -0.3614%; 1.08: 37.64%; 1.09: -7.336%; 1.10: -26.93%; 1.11: 17.08%; 1.12: 21.23%; 1.13: 41.15%
| Z-Score (Absolute) |  | 0.39 | 0.08 | 2.45 | 0.03 | 0.11 | 0.03 | 4.98 | 1.29 | 0.2 | 0.85 | 0.61 | 0.74 | 1.4 | 4.98
| Analytical results for Lu in OREAS 627 (Certified Value 0.21 ppm) | [24 empty] | Certified
Lu | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab |  |  | 1
| No | No | 1.01: 1.09; 1.02: 1.10; 1.03: 1.13; 1.04: 1.19; 1.05: 1.20; 1.06: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*MS; 1.03: PF*OES/MS; 1.04: BF*MS; 1.05: PF*MS; 1.06: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 0.19; 1.02: 0.2; 1.03: 0.20213; 1.04: 0.225; 1.05: < 2; 1.06: 0.22 |  |  | 1
| 1 | 2 | 1.01: 0.19; 1.02: 0.2; 1.03: 0.20185; 1.04: 0.223; 1.05: < 2; 1.06: 0.22 |  |  | 35
| 1 | 3 | 1.01: 0.19; 1.02: 0.2; 1.03: 0.212; 1.04: 0.224; 1.05: < 2; 1.06: 0.22 |  |  | 16
| 1 | 4 | 1.01: 0.19; 1.02: 0.2; 1.03: 0.19385; 1.04: 0.222; 1.05: < 2; 1.06: 0.23 |  |  | 0.20998
| 1 | 5 | 1.01: 0.19; 1.02: 0.2; 1.03: 0.22924; 1.04: 0.226; 1.05: < 2; 1.06: 0.23 |  |  | 142
| 1 | 6 | 1.01: 0.2; 1.02: 0.2; 1.03: 0.22223; 1.04: 0.228; 1.05: < 2; 1.06: 0.22
| Mean |  | 1.01: 0.19167; 1.02: 0.2; 1.03: 0.21022; 1.04: 0.22467; 1.06: 0.22333
| Median |  | 1.01: 0.19; 1.02: 0.2; 1.03: 0.20707; 1.04: 0.2245; 1.06: 0.22
| Std Dev. |  | 1.01: 0.0040825; 1.02: 3.0405e-17; 1.03: 0.01351; 1.04: 0.0021602; 1.06: 0.005164
| Rel.Std.Dev. |  | 1.01: 2.13%; 1.02: 1.52e-14%; 1.03: 6.427%; 1.04: 0.9615%; 1.06: 2.312%
| PDM3 |  | 1.01: -8.72%; 1.02: -4.751%; 1.03: 0.1139%; 1.04: 6.996%; 1.06: 6.361%
| Z-Score (Absolute) |  | 1.01: 1.37; 1.02: 0.92; 1.03: 0.36; 1.04: 0.43; 1.05: 42.82; 1.06: 0.36
| Analytical results for Mg in OREAS 627 (Certified Value 0.396 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES |  | 3
3
| 1 | 1 | 0.52 | 0.39 | 0.42 | 0.38 | 0.398 | 0.338 | 0.398 | 0.368 | 0.39 | 0.41 | 0.4 | 0.4 | 0.39172 | 0.43524 | 0.41 | 0.38 | 0.4168 | 0.50688 | 0.39 | 0.4 | 0.36 | 0.42 |  | 1
| 1 | 2 | 0.55 | 0.39 | 0.43 | 0.38 | 0.404 | 0.338 | 0.386 | 0.374 | 0.39 | 0.4 | 0.42 | 0.39 | 0.42831 | 0.43692 | 0.42 | 0.38 | 0.4176 | 0.50535 | 0.4 | 0.41 | 0.37 | 0.42 |  | 0
| 1 | 3 | 0.53 | 0.39 | 0.43 | 0.386 | 0.398 | 0.362 | 0.398 | 0.38 | 0.4 | 0.43 | 0.4 | 0.39 | 0.39857 | 0.44901 | 0.42 | 0.38 | 0.4098 | 0.50945 | 0.41 | 0.39 | 0.36 | 0.42 |  | 16
| 1 | 4 | 0.58 | 0.38 | 0.42 | 0.38 | 0.38 | 0.362 | 0.392 | 0.374 | 0.39 | 0.4 | 0.45 | 0.38 | 0.40766 | 0.41753 | 0.42 | 0.38 | 0.409 | 0.50788 | 0.39 | 0.4 | 0.36 | 0.4 |  | 0.39593
| 1 | 5 | 0.56 | 0.39 | 0.44 | 0.386 | 0.386 | 0.362 | 0.386 | 0.374 | 0.4 | 0.39 | 0.41 | 0.4 | 0.41205 | 0.41585 | 0.43 | 0.39 | 0.4121 | 0.49584 | 0.4 | 0.39 | 0.36 | 0.42 |  | 143
| 1 | 6 | 0.58 | 0.4 | 0.41 | 0.404 | 0.386 | 0.362 | 0.386 | 0.368 | 0.4 | 0.36 | 0.41 | 0.39 | 0.39836 | 0.4071 | 0.42 | 0.38 | 0.3984 | 0.50527 | 0.39 | 0.4 | 0.36 | 0.42
| Mean |  | 0.55333 | 0.39 | 0.425 | 0.386 | 0.392 | 0.354 | 0.391 | 0.373 | 0.395 | 0.39833 | 0.415 | 0.39167 | 0.40611 | 0.42694 | 0.42 | 0.38167 | 0.41062 | 0.50511 | 0.39667 | 0.39833 | 0.36167 | 0.41667
| Median |  | 0.555 | 0.39 | 0.425 | 0.383 | 0.392 | 0.362 | 0.389 | 0.374 | 0.395 | 0.4 | 0.41 | 0.39 | 0.40312 | 0.42639 | 0.42 | 0.38 | 0.41095 | 0.50611 | 0.395 | 0.4 | 0.36 | 0.42
| Std Dev. |  | 0.025033 | 0.0063246 | 0.010488 | 0.0092952 | 0.0092952 | 0.012394 | 0.0058992 | 0.0045166 | 0.0054772 | 0.023166 | 0.018708 | 0.0075277 | 0.013072 | 0.015882 | 0.0063246 | 0.0040825 | 0.0069531 | 0.0048104 | 0.008165 | 0.0075277 | 0.0040825 | 0.008165
| Rel.Std.Dev. |  | 4.524% | 1.622% | 2.468% | 2.408% | 2.371% | 3.501% | 1.509% | 1.211% | 1.387% | 5.816% | 4.508% | 1.922% | 3.219% | 3.72% | 1.506% | 1.07% | 1.693% | 0.9524% | 2.058% | 1.89% | 1.129% | 1.96%
| PDM3 |  | 39.75% | -1.499% | 7.341% | -2.509% | -0.9936% | -10.59% | -1.246% | -5.792% | -0.2359% | 0.606% | 4.815% | -1.078% | 2.571% | 7.832% | 6.078% | -3.603% | 3.708% | 27.57% | 0.185% | 0.606% | -8.655% | 5.236%
| Z-Score (Absolute) |  | 7.26 | 0.35 | 1.28 | 0.54 | 0.26 | 2.04 | 0.3 | 1.15 | 0.12 | 0.04 | 0.82 | 0.27 | 0.4 | 1.37 | 1.05 | 0.74 | 0.61 | 5.01 | 0.04 | 0.04 | 1.67 | 0.89
| Analytical results for Mn in OREAS 627 (Certified Value 0.711 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 3
3
| 1 | 1 | 0.7046 | 0.744 | 0.7328 | 0.697 | 0.72 | 0.705 | 0.705 | 0.689 | 0.721 | 0.7488 | 0.73 | 0.6659 | 0.69139 | 0.81603 | 0.6823 | 0.714 | 0.7145 | 0.7014 | 0.7391 | 0.64 | 0.73 | 0.72 |  | 1
| 1 | 2 | 0.7065 | 0.745 | 0.7424 | 0.713 | 0.728 | 0.705 | 0.697 | 0.689 | 0.723 | 0.7393 | 0.71 | 0.6526 | 0.68331 | 0.81711 | 0.6958 | 0.713 | 0.7181 | 0.7048 | 0.7305 | 0.636 | 0.74 | 0.71 |  | 21
| 1 | 3 | 0.7063 | 0.728 | 0.7257 | 0.697 | 0.72 | 0.705 | 0.697 | 0.713 | 0.727 | 0.7303 | 0.71 | 0.6584 | 0.69276 | 0.81265 | 0.6949 | 0.711 | 0.7082 | 0.7047 | 0.7334 | 0.633 | 0.74 | 0.72 |  | 16
| 1 | 4 | 0.704 | 0.73 | 0.7196 | 0.713 | 0.705 | 0.728 | 0.713 | 0.689 | 0.723 | 0.7415 | 0.73 | 0.6404 | 0.69334 | 0.81649 | 0.6832 | 0.706 | 0.7142 | 0.6983 | 0.7332 | 0.634 | 0.72 | 0.71 |  | 0.71122
| 1 | 5 | 0.7002 | 0.722 | 0.7443 | 0.689 | 0.713 | 0.713 | 0.697 | 0.689 | 0.734 | 0.7173 | 0.71 | 0.6698 | 0.70183 | 0.80737 | 0.6963 | 0.722 | 0.7156 | 0.704 | 0.7319 | 0.642 | 0.74 | 0.71 |  | 144
| 1 | 6 | 0.706 | 0.743 | 0.7301 | 0.689 | 0.728 | 0.713 | 0.697 | 0.689 | 0.728 | 0.7367 | 0.71 | 0.6499 | 0.70294 | 0.80851 | 0.6945 | 0.712 | 0.7056 | 0.7059 | 0.7326 | 0.64 | 0.73 | 0.71
| Mean |  | 0.7046 | 0.73533 | 0.73248 | 0.69967 | 0.719 | 0.7115 | 0.701 | 0.693 | 0.726 | 0.73565 | 0.71667 | 0.65617 | 0.69426 | 0.81302 | 0.69117 | 0.713 | 0.7127 | 0.70318 | 0.73345 | 0.6375 | 0.73333 | 0.71333
| Median |  | 0.7053 | 0.7365 | 0.73145 | 0.697 | 0.72 | 0.709 | 0.697 | 0.689 | 0.725 | 0.738 | 0.71 | 0.6555 | 0.69305 | 0.81434 | 0.6947 | 0.7125 | 0.71435 | 0.70435 | 0.7329 | 0.638 | 0.735 | 0.71
| Std Dev. |  | 0.002374 | 0.0098725 | 0.0095472 | 0.01093 | 0.0088994 | 0.0089833 | 0.0066933 | 0.009798 | 0.0047329 | 0.010834 | 0.010328 | 0.010827 | 0.0072657 | 0.0042486 | 0.0065567 | 0.0052154 | 0.0047691 | 0.002828 | 0.0029602 | 0.0036742 | 0.008165 | 0.005164
| Rel.Std.Dev. |  | 0.3369% | 1.343% | 1.303% | 1.562% | 1.238% | 1.263% | 0.9548% | 1.414% | 0.6519% | 1.473% | 1.441% | 1.65% | 1.047% | 0.5226% | 0.9486% | 0.7315% | 0.6692% | 0.4022% | 0.4036% | 0.5764% | 1.113% | 0.7239%
| PDM3 |  | -0.9307% | 3.391% | 2.99% | -1.624% | 1.094% | 0.03945% | -1.437% | -2.562% | 2.078% | 3.435% | 0.7659% | -7.741% | -2.384% | 14.31% | -2.819% | 0.2504% | 0.2082% | -1.13% | 3.126% | -10.37% | 3.109% | 0.2972%
| Z-Score (Absolute) |  | 0.35 | 0.95 | 0.83 | 0.56 | 0.26 | 0.07 | 0.51 | 0.83 | 0.56 | 0.97 | 0.16 | 2.4 | 0.79 | 4.24 | 0.92 | 0.01 | 0.01 | 0.41 | 0.87 | 3.19 | 0.87 | 0.02
| Analytical results for Mo in OREAS 627 (Certified Value 32.5 ppm) | [24 empty] | Certified
Mo | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab; 1.13: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.12; 1.06: 1.13; 1.07: 1.15; 1.08: 1.16; 1.09: 1.17; 1.10: 1.20; 1.11: 1.23; 1.12: 1.25; 1.13: 1.26 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*OES; 1.09: PF*OES/MS; 1.10: PF*MS; 1.11: PF*MS; 1.12: PF*OES; 1.13: PF*OES |  | 1
2
| 1 | 1 | 1.01: 30; 1.02: 33; 1.03: 27; 1.04: 33; 1.05: 33; 1.06: 30.3882; 1.07: 37.6496; 1.08: 31.6; 1.09: 37; 1.10: 38; 1.11: 31; 1.12: 30; 1.13: 40 |  | 1
| 1 | 2 | 1.01: 22; 1.02: 28; 1.03: 29; 1.04: 33; 1.05: 32; 1.06: 31.3532; 1.07: 36.5616; 1.08: 31.8; 1.09: 36; 1.10: 33; 1.11: 31; 1.12: 30; 1.13: 40 |  | 37
| 1 | 3 | 1.01: 28; 1.02: 29; 1.03: 29; 1.04: 34; 1.05: 34; 1.06: 32.8204; 1.07: 37.5544; 1.08: 31; 1.09: 37; 1.10: 32; 1.11: 33; 1.12: 29; 1.13: 30 |  | 16
| 1 | 4 | 1.01: 27; 1.02: 29; 1.03: 30; 1.04: 33; 1.05: 30; 1.06: 32.6491; 1.07: 38.47; 1.08: 32; 1.09: 37; 1.10: 33; 1.11: 32; 1.12: 31; 1.13: 40 |  | 32.4602
| 1 | 5 | 1.01: 22; 1.02: 27; 1.03: 28; 1.04: 33; 1.05: 31; 1.06: 32.0812; 1.07: 37.9891; 1.08: 31; 1.09: 37; 1.10: 34; 1.11: 33; 1.12: 31; 1.13: 40 |  | 145
| 1 | 6 | 1.01: 24; 1.02: 30; 1.03: 30; 1.04: 34; 1.05: 31; 1.06: 33.1056; 1.07: 36.551; 1.08: 31.6; 1.09: 38; 1.10: 34; 1.11: 33; 1.12: 31; 1.13: 30
| Mean |  | 1.01: 25.5; 1.02: 29.3333; 1.03: 28.8333; 1.04: 33.3333; 1.05: 31.8333; 1.06: 32.0663; 1.07: 37.4626; 1.08: 31.5; 1.09: 37; 1.10: 34; 1.11: 32.1667; 1.12: 30.3333; 1.13: 36.6667
| Median |  | 1.01: 25.5; 1.02: 29; 1.03: 29; 1.04: 33; 1.05: 31.5; 1.06: 32.3651; 1.07: 37.602; 1.08: 31.6; 1.09: 37; 1.10: 33.5; 1.11: 32.5; 1.12: 30.5; 1.13: 40
| Std Dev. |  | 1.01: 3.3317; 1.02: 2.0656; 1.03: 1.169; 1.04: 0.5164; 1.05: 1.472; 1.06: 1.0308; 1.07: 0.77181; 1.08: 0.41473; 1.09: 0.63246; 1.10: 2.0976; 1.11: 0.98319; 1.12: 0.8165; 1.13: 5.164
| Rel.Std.Dev. |  | 1.01: 13.07%; 1.02: 7.042%; 1.03: 4.054%; 1.04: 1.549%; 1.05: 4.624%; 1.06: 3.215%; 1.07: 2.06%; 1.08: 1.317%; 1.09: 1.709%; 1.10: 6.169%; 1.11: 3.057%; 1.12: 2.692%; 1.13: 14.08%
| PDM3 |  | 1.01: -21.44%; 1.02: -9.633%; 1.03: -11.17%; 1.04: 2.69%; 1.05: -1.931%; 1.06: -1.214%; 1.07: 15.41%; 1.08: -2.958%; 1.09: 13.99%; 1.10: 4.744%; 1.11: -0.9043%; 1.12: -6.552%; 1.13: 12.96%
| Z-Score (Absolute) |  | 1.01: 2.37; 1.02: 0.96; 1.03: 1.15; 1.04: 0.51; 1.05: 0.04; 1.06: 0.04; 1.07: 2.03; 1.08: 0.17; 1.09: 1.86; 1.10: 0.75; 1.11: 0.08; 1.12: 0.59; 1.13: NA
| Analytical results for Na in OREAS 627 (Indicative Value 1.08 wt.%) | [24 empty] | Indicative
Na | Round | Replicate | 1.01: Lab |  | 1
| No | No | 1.01: 1.16 |  | wt.%
|  |  | 1.01: PF*OES |  | 2
2
| 1 | 1 | 1.01: 1.07 |  | 1
| 1 | 2 | 1.01: 1.09 |  | 13
| 1 | 3 | 1.01: 1.11 |  | 16
| 1 | 4 | 1.01: 1.06 |  | 1.0817
| 1 | 5 | 1.01: 1.07 |  | 19
| 1 | 6 | 1.01: 1.09
| Mean |  | 1.01: 1.0817
| Median |  | 1.01: 1.08
| Std Dev. |  | 1.01: 0.018348
| Rel.Std.Dev. |  | 1.01: 1.696%
| PDM3 |  | 1.01: -3.109e-13%
Z-Score (Absolute) | bar: 1.01
| Analytical results for Nb in OREAS 627 (Certified Value 11.2 ppm) | [24 empty] | Certified
Nb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.20; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 1
1
| 1 | 1 | 1.01: 9; 1.02: 13.2; 1.03: 6; 1.04: 12; 1.05: 12.5117; 1.06: 11.6868; 1.07: 11; 1.08: 10 |  | 1
| 1 | 2 | 1.01: 8; 1.02: 12.8; 1.03: 9; 1.04: 10; 1.05: 12.1276; 1.06: 11.6589; 1.07: 12; 1.08: 11 |  | 38
| 1 | 3 | 1.01: 8; 1.02: 12.2; 1.03: 9; 1.04: < 10; 1.05: 11.7397; 1.06: 11.5357; 1.07: 11; 1.08: 11 |  | 16
| 1 | 4 | 1.01: 9; 1.02: 12.7; 1.03: 10; 1.04: 10; 1.05: 11.7523; 1.06: 11.5812; 1.07: 12; 1.08: 11 |  | 11.1891
| 1 | 5 | 1.01: 8; 1.02: 11.9; 1.03: 8; 1.04: < 10; 1.05: 11.9783; 1.06: 11.745; 1.07: 12; 1.08: 12 |  | 146
| 1 | 6 | 1.01: 8; 1.02: 11.8; 1.03: 9; 1.04: 11; 1.05: 11.9365; 1.06: 11.5893; 1.07: 11; 1.08: 11
| Mean |  | 1.01: 8.3333; 1.02: 12.4333; 1.03: 8.5; 1.04: 10.75; 1.05: 12.0077; 1.06: 11.6328; 1.07: 11.5; 1.08: 11
| Median |  | 1.01: 8; 1.02: 12.45; 1.03: 9; 1.04: 10.5; 1.05: 11.9574; 1.06: 11.6241; 1.07: 11.5; 1.08: 11
| Std Dev. |  | 1.01: 0.5164; 1.02: 0.55377; 1.03: 1.3784; 1.04: 0.95743; 1.05: 0.28683; 1.06: 0.077685; 1.07: 0.54772; 1.08: 0.63246
| Rel.Std.Dev. |  | 1.01: 6.197%; 1.02: 4.454%; 1.03: 16.22%; 1.04: 8.906%; 1.05: 2.389%; 1.06: 0.6678%; 1.07: 4.763%; 1.08: 5.75%
| PDM3 |  | 1.01: -25.52%; 1.02: 11.12%; 1.03: -24.03%; 1.04: -3.925%; 1.05: 7.316%; 1.06: 3.966%; 1.07: 2.778%; 1.08: -1.69%
| Z-Score (Absolute) |  | 1.01: 2.03; 1.02: 0.82; 1.03: 1.91; 1.04: 1.68; 1.05: 0.53; 1.06: 0.27; 1.07: 0.17; 1.08: 0.17
| Analytical results for Nd in OREAS 627 (Certified Value 30.1 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  | 1
2
| 1 | 1 | 1.01: 31.1; 1.02: 27; 1.03: 31.2; 1.04: 30.9593; 1.05: 32.292; 1.06: 35; 1.07: 27.2188; 1.08: 28.8 |  | 1
| 1 | 2 | 1.01: 31.6; 1.02: 28; 1.03: 30.8; 1.04: 28.9393; 1.05: 32.543; 1.06: 29; 1.07: 27.4501; 1.08: 29.1 |  | 7
| 1 | 3 | 1.01: 34.4; 1.02: 29; 1.03: 30.5; 1.04: 30.0185; 1.05: 32.651; 1.06: 29; 1.07: 27.7762; 1.08: 30.3 |  | 16
| 1 | 4 | 1.01: 29.4; 1.02: 29; 1.03: 30.8; 1.04: 29.4356; 1.05: 32.446; 1.06: 31; 1.07: 27.6044; 1.08: 29.9 |  | 30.0793
| 1 | 5 | 1.01: 31.3; 1.02: 28; 1.03: 29.4; 1.04: 30.7773; 1.05: 32.604; 1.06: 33; 1.07: 27.2342; 1.08: 30.6 |  | 147
| 1 | 6 | 1.01: 30; 1.02: 29; 1.03: 30; 1.04: 29.1233; 1.05: 32.582; 1.06: 29; 1.07: 27.151; 1.08: 29.8
| Mean |  | 1.01: 31.3; 1.02: 28.3333; 1.03: 30.45; 1.04: 29.8755; 1.05: 32.5197; 1.06: 31; 1.07: 27.4058; 1.08: 29.75
| Median |  | 1.01: 31.2; 1.02: 28.5; 1.03: 30.65; 1.04: 29.7271; 1.05: 32.5625; 1.06: 30; 1.07: 27.3421; 1.08: 29.85
| Std Dev. |  | 1.01: 1.7344; 1.02: 0.8165; 1.03: 0.65038; 1.04: 0.8537; 1.05: 0.13119; 1.06: 2.5298; 1.07: 0.24817; 1.08: 0.6892
| Rel.Std.Dev. |  | 1.01: 5.541%; 1.02: 2.882%; 1.03: 2.136%; 1.04: 2.858%; 1.05: 0.4034%; 1.06: 8.161%; 1.07: 0.9055%; 1.08: 2.317%
| PDM3 |  | 1.01: 4.058%; 1.02: -5.805%; 1.03: 1.232%; 1.04: -0.6774%; 1.05: 8.113%; 1.06: 3.061%; 1.07: -8.888%; 1.08: -1.095%
| Z-Score (Absolute) |  | 1.01: 0.78; 1.02: 1.25; 1.03: 0.2; 1.04: 0.2; 1.05: 1.61; 1.06: 0.57; 1.07: 1.88; 1.08: 0.28
| Analytical results for Ni in OREAS 627 (Indicative Value 17 ppm) | [24 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 1
1
| 1 | 1 | 15 | 20 | 77 | < 20 | < 20 | < 20 | 20 | < 20 | < 10 | < 20 | < 50 | < 20 | 35.4099 | < 20 | 8.88 | < 20 | NA | 6 | < 10 | 10 | < 50 | < 100 |  | 1
| 1 | 2 | 14 | 20 | 57 | < 20 | < 20 | < 20 | < 20 | < 20 | 12 | < 20 | < 50 | < 20 | 28.7932 | < 20 | 8.39 | 20 | NA | 8 | < 10 | 20 | < 50 | < 100 |  | 14
| 1 | 3 | 14 | 20 | 63 | < 20 | < 20 | < 20 | 20 | < 20 | 10 | < 20 | < 50 | < 20 | 31.4413 | < 20 | 8.6 | < 20 | NA | 9 | < 10 | 10 | < 50 | < 100 |  | 16
| 1 | 4 | 13 | 20 | 52 | < 20 | < 20 | < 20 | < 20 | < 20 | < 10 | < 20 | < 50 | < 20 | 24.5432 | < 20 | 8.01 | < 20 | NA | 8 | < 10 | 30 | < 50 | < 100 |  | 17.0402
| 1 | 5 | 14 | 20 | 117 | 20 | < 20 | < 20 | < 20 | < 20 | 11 | < 20 | < 50 | < 20 | 37.9733 | < 20 | 9.11 | < 20 | NA | 7 | < 10 | 20 | < 50 | < 100 |  | 20
| 1 | 6 | 14 | 20 | < 50 | < 20 | < 20 | < 20 | < 20 | < 20 | 12 | < 20 | < 50 | < 20 | 26.3232 | < 20 | 8.44 | < 20 | NA | 7 | < 10 | 20 | < 50 | < 100
| Mean |  | 1.01: 14; 1.02: 20; 1.03: 73.2; 1.04: 20; 1.07: 20; 1.09: 11.25; 1.13: 30.7474; 1.16: 8.5717; 1.17: 20; 1.23: 7.5; 1.26: 18.3333
| Median |  | 1.01: 14; 1.02: 20; 1.03: 63; 1.04: 20; 1.07: 20; 1.09: 11.5; 1.13: 30.1173; 1.16: 8.52; 1.17: 20; 1.23: 7.5; 1.26: 20
| Std Dev. |  | 1.01: 0.63246; 1.02: 0; 1.03: 26.2145; 1.07: 0; 1.09: 0.95743; 1.13: 5.2214; 1.16: 0.38768; 1.23: 1.0488; 1.26: 7.5277
| Rel.Std.Dev. |  | 1.01: 4.518%; 1.02: 0%; 1.03: 35.81%; 1.07: 0%; 1.09: 8.51%; 1.13: 16.98%; 1.16: 4.523%; 1.23: 13.98%; 1.26: 41.06%
| PDM3 |  | 1.01: -17.84%; 1.02: 17.37%; 1.03: 329.6%; 1.04: 17.37%; 1.07: 17.37%; 1.09: -33.98%; 1.13: 80.44%; 1.16: -49.7%; 1.17: 17.37%; 1.23: -55.99%; 1.26: 7.588%
| Z-Score (Absolute) |  | 0.63 | 2.25 | 14.43 | 0 | 0.45 | 0.45 | 0.45 | 0.45 | 0.67 | 0.45 | 3.6 | 0.45 | 5.15 | 0.45 | 0.83 | 0 | NA | 1.12 | 1.8 | 1.8 | 3.6 | 10.34
| Analytical results for P in OREAS 627 (Certified Value 0.038 wt.%) | [24 empty] | Certified
P | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  | 1
| No | No | 1.01: 1.01; 1.02: 1.09; 1.03: 1.10; 1.04: 1.12; 1.05: 1.13; 1.06: 1.16; 1.07: 1.17; 1.08: 1.20; 1.09: 1.23; 1.10: 1.25; 1.11: 1.26 |  | wt.%
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES; 1.05: PF*OES/MS; 1.06: PF*OES; 1.07: PF*OES/MS; 1.08: PF*MS; 1.09: PF*OES; 1.10: PF*OES; 1.11: PF*OES |  | 3
3
| 1 | 1 | 1.01: 0.0251; 1.02: 0.06; 1.03: 0.04; 1.04: 0.04; 1.05: < 0.05; 1.06: 0.04; 1.07: 0.04; 1.08: 0.0436; 1.09: 0.03; 1.10: 0.05; 1.11: 0.03 |  | 1
| 1 | 2 | 1.01: 0.0256; 1.02: 0.06; 1.03: 0.04; 1.04: 0.04; 1.05: < 0.05; 1.06: 0.04; 1.07: 0.04; 1.08: 0.0405; 1.09: 0.03; 1.10: 0.05; 1.11: 0.03 |  | 24
| 1 | 3 | 1.01: 0.0253; 1.02: 0.06; 1.03: 0.04; 1.04: 0.04; 1.05: < 0.05; 1.06: 0.04; 1.07: 0.04; 1.08: 0.0433; 1.09: 0.04; 1.10: 0.06; 1.11: 0.03 |  | 16
| 1 | 4 | 1.01: 0.0261; 1.02: 0.06; 1.03: 0.04; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.04; 1.07: 0.04; 1.08: 0.0417; 1.09: 0.03; 1.10: 0.05; 1.11: 0.04 |  | 0.037643
| 1 | 5 | 1.01: 0.0253; 1.02: 0.06; 1.03: 0.04; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.04; 1.07: 0.04; 1.08: 0.0437; 1.09: 0.03; 1.10: 0.06; 1.11: 0.04 |  | 148
| 1 | 6 | 1.01: 0.0255; 1.02: 0.06; 1.03: 0.03; 1.04: 0.03; 1.05: < 0.05; 1.06: 0.04; 1.07: 0.04; 1.08: 0.037; 1.09: 0.03; 1.10: 0.05; 1.11: 0.03
| Mean |  | 1.01: 0.025483; 1.02: 0.06; 1.03: 0.038333; 1.04: 0.035; 1.06: 0.04; 1.07: 0.04; 1.08: 0.041633; 1.09: 0.031667; 1.10: 0.053333; 1.11: 0.033333
| Median |  | 1.01: 0.0254; 1.02: 0.06; 1.03: 0.04; 1.04: 0.035; 1.06: 0.04; 1.07: 0.04; 1.08: 0.0425; 1.09: 0.03; 1.10: 0.05; 1.11: 0.03
| Std Dev. |  | 1.01: 0.00034881; 1.02: 0; 1.03: 0.0040825; 1.04: 0.0054772; 1.06: 0; 1.07: 0; 1.08: 0.0025951; 1.09: 0.0040825; 1.10: 0.005164; 1.11: 0.005164
| Rel.Std.Dev. |  | 1.01: 1.369%; 1.02: 0%; 1.03: 10.65%; 1.04: 15.65%; 1.06: 0%; 1.07: 0%; 1.08: 6.233%; 1.09: 12.89%; 1.10: 9.682%; 1.11: 15.49%
| PDM3 |  | 1.01: -32.3%; 1.02: 59.39%; 1.03: 1.835%; 1.04: -7.02%; 1.06: 6.263%; 1.07: 6.263%; 1.08: 10.6%; 1.09: -15.88%; 1.10: 41.68%; 1.11: -11.45%
| Z-Score (Absolute) |  | 1.01: 1.73; 1.02: 2.92; 1.03: 0; 1.04: 0.45; 1.05: 1.8; 1.06: 0.22; 1.07: 0.22; 1.08: 0.45; 1.09: 0.9; 1.10: 2.02; 1.11: 0.67
| Analytical results for Pb in OREAS 627 (Certified Value 2.21 wt.%) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | 2
3
| 1 | 1 | 2.13 | 2.22 | 2.1041 | 2.08 | 2.16 | 2.19 | 2.13 | 2.1 | 2.04 | 2.354 | 2.2 | 2.2253 | 2.2769 | > 1 | 2.13 | 2.2099 | 2.3458 | 2.2902 | 2.24 | 2.3186 | 2.182 | 2.24 | 2.23 | 1
| 1 | 2 | 2.12 | 2.25 | 2.1186 | 2.11 | 2.19 | 2.21 | 2.14 | 2.1 | 2.16 | 2.372 | 2.21 | 2.1781 | 2.2327 | > 1 | 2.2 | 2.2015 | 2.343 | 2.2943 | 2.24 | 2.3063 | 2.163 | 2.27 | 2.24 | 25
| 1 | 3 | 2.14 | 2.23 | 2.0679 | 2.14 | 2.18 | 2.19 | 2.14 | 2.15 | 2.18 | 2.3757 | 2.17 | 2.1975 | 2.2655 | > 1 | 2.16 | 2.2039 | 2.2643 | 2.2617 | 2.23 | 2.32 | 2.176 | 2.27 | 2.23 | 16
| 1 | 4 | 2.15 | 2.24 | 2.0578 | 2.13 | 2.14 | 2.23 | 2.14 | 2.1 | 2.2 | 2.381 | 2.26 | 2.1389 | 2.2795 | > 1 | 2.15 | 2.2084 | 2.3381 | 2.2249 | 2.21 | 2.3287 | 2.189 | 2.16 | 2.23 | 2.2071
| 1 | 5 | 2.14 | 2.24 | 2.1294 | 2.11 | 2.2 | 2.22 | 2.13 | 2.1 | 2.1 | 2.3299 | 2.21 | 2.2465 | 2.2478 | > 1 | 2.16 | 2.2073 | 2.3745 | 2.2365 | 2.24 | 2.3897 | 2.17 | 2.24 | 2.21 | 149
| 1 | 6 | 2.15 | 2.23 | 2.0681 | 2.11 | 2.21 | 2.2 | 2.16 | 2.1 | 2.24 | 2.3241 | 2.19 | 2.169 | 2.2197 | > 1 | 2.15 | 2.2068 | 2.3169 | 2.2094 | 2.23 | 2.3762 | 2.142 | 2.25 | 2.21
| Mean |  | 2.1383 | 2.235 | 2.091 | 2.1133 | 2.18 | 2.2067 | 2.14 | 2.1083 | 2.1533 | 2.3561 | 2.2067 | 2.1926 | 2.2537 |  | 2.1583 | 2.2063 | 2.3304 | 2.2528 | 2.2317 | 2.3399 | 2.1703 | 2.2383 | 2.225
| Median |  | 2.14 | 2.235 | 2.0861 | 2.11 | 2.185 | 2.205 | 2.14 | 2.1 | 2.17 | 2.363 | 2.205 | 2.1878 | 2.2567 |  | 2.155 | 2.2071 | 2.3405 | 2.2491 | 2.235 | 2.3243 | 2.173 | 2.245 | 2.23
| Std Dev. |  | 0.01169 | 0.010488 | 0.030226 | 0.020656 | 0.026077 | 0.01633 | 0.010954 | 0.020412 | 0.072296 | 0.02438 | 0.030111 | 0.039114 | 0.024391 |  | 0.023166 | 0.0030796 | 0.037287 | 0.03502 | 0.01169 | 0.034356 | 0.016573 | 0.040702 | 0.012247
| Rel.Std.Dev. |  | 0.5467% | 0.4693% | 1.446% | 0.9774% | 1.196% | 0.74% | 0.5119% | 0.9682% | 3.357% | 1.035% | 1.365% | 1.784% | 1.082% |  | 1.073% | 0.1396% | 1.6% | 1.554% | 0.5238% | 1.468% | 0.7636% | 1.818% | 0.5504%
| PDM3 |  | -3.117% | 1.262% | -5.263% | -4.25% | -1.23% | -0.02139% | -3.042% | -4.477% | -2.438% | 6.75% | -0.02139% | -0.661% | 2.109% |  | -2.211% | -0.038% | 5.586% | 2.071% | 1.111% | 6.016% | -1.668% | 1.413% | 0.8092%
| Z-Score (Absolute) |  | 0.98 | 0.41 | 1.66 | 1.34 | 0.38 | 0 | 0.96 | 1.41 | 0.77 | 2.16 | 0 | 0.2 | 0.68 | NA | 0.69 | 0 | 1.79 | 0.67 | 0.36 | 1.92 | 0.52 | 0.46 | 0.27
| Analytical results for Pr in OREAS 627 (Certified Value 8.27 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | 2
3
| 1 | 1 | 1.01: 8.3; 1.02: 6.83; 1.03: 8.2; 1.04: 8.2346; 1.05: 8.464; 1.06: 10; 1.07: 8.6043; 1.08: 7.83 | 1
| 1 | 2 | 1.01: 8; 1.02: 7.16; 1.03: 8.1; 1.04: 7.9226; 1.05: 8.475; 1.06: 9; 1.07: 8.5953; 1.08: 7.52 | 9
| 1 | 3 | 1.01: 8.5; 1.02: 7.24; 1.03: 8.5; 1.04: 8.3474; 1.05: 8.504; 1.06: 9; 1.07: 8.932; 1.08: 8.04 | 16
| 1 | 4 | 1.01: 7.1; 1.02: 7.43; 1.03: 8.4; 1.04: 8.1597; 1.05: 8.632; 1.06: 9; 1.07: 8.5167; 1.08: 7.93 | 8.2696
| 1 | 5 | 1.01: 8.4; 1.02: 6.94; 1.03: 8.2; 1.04: 8.545; 1.05: 8.457; 1.06: 10; 1.07: 8.6136; 1.08: 8.23 | 150
| 1 | 6 | 1.01: 7.5; 1.02: 7.28; 1.03: 8.2; 1.04: 8.1234; 1.05: 8.501; 1.06: 9; 1.07: 8.4356; 1.08: 8.03
| Mean |  | 1.01: 7.9667; 1.02: 7.1467; 1.03: 8.2667; 1.04: 8.2221; 1.05: 8.5055; 1.06: 9.3333; 1.07: 8.6162; 1.08: 7.93
| Median |  | 1.01: 8.15; 1.02: 7.2; 1.03: 8.2; 1.04: 8.1972; 1.05: 8.488; 1.06: 9; 1.07: 8.5998; 1.08: 7.98
| Std Dev. |  | 1.01: 0.55737; 1.02: 0.22358; 1.03: 0.15055; 1.04: 0.21142; 1.05: 0.064853; 1.06: 0.5164; 1.07: 0.169; 1.08: 0.24091
| Rel.Std.Dev. |  | 1.01: 6.996%; 1.02: 3.128%; 1.03: 1.821%; 1.04: 2.571%; 1.05: 0.7625%; 1.06: 5.533%; 1.07: 1.961%; 1.08: 3.038%
| PDM3 |  | 1.01: -3.663%; 1.02: -13.58%; 1.03: -0.03505%; 1.04: -0.5736%; 1.05: 2.853%; 1.06: 12.86%; 1.07: 4.192%; 1.08: -4.106%
| Z-Score (Absolute) |  | 1.01: 0.61; 1.02: 2.56; 1.03: 0.11; 1.04: 0; 1.05: 0.67; 1.06: NA; 1.07: 0.94; 1.08: 0.7
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Rb in OREAS 627 (Certified Value 127 ppm) | [24 empty] | Certified
Rb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.15; 1.06: 1.17; 1.07: 1.20; 1.08: 1.23 | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: PF*MS; 1.08: PF*MS | 0
0
| 1 | 1 | 1.01: 130; 1.02: 108; 1.03: 131; 1.04: 125.0119; 1.05: 134.7896; 1.06: 131; 1.07: 126; 1.08: 117 | 1
| 1 | 2 | 1.01: 128; 1.02: 114; 1.03: 131.1; 1.04: 122.8783; 1.05: 134.3584; 1.06: 133; 1.07: 127; 1.08: 118 | 43
| 1 | 3 | 1.01: 130; 1.02: 115; 1.03: 132.5; 1.04: 125.8354; 1.05: 139.8062; 1.06: 136; 1.07: 124; 1.08: 121 | 16
| 1 | 4 | 1.01: 129; 1.02: 117; 1.03: 131.6; 1.04: 126.6305; 1.05: 133.2111; 1.06: 132; 1.07: 127; 1.08: 117 | 126.9695
| 1 | 5 | 1.01: 123; 1.02: 110; 1.03: 129.4; 1.04: 127.7941; 1.05: 139.906; 1.06: 138; 1.07: 128; 1.08: 122 | 151
| 1 | 6 | 1.01: 130; 1.02: 119; 1.03: 131.1; 1.04: 125.0268; 1.05: 135.1891; 1.06: 134; 1.07: 126; 1.08: 121
| Mean |  | 1.01: 128.3333; 1.02: 113.8333; 1.03: 131.1167; 1.04: 125.5295; 1.05: 136.2101; 1.06: 134; 1.07: 126.3333; 1.08: 119.3333
| Median |  | 1.01: 129.5; 1.02: 114.5; 1.03: 131.1; 1.04: 125.4311; 1.05: 134.9894; 1.06: 133.5; 1.07: 126.5; 1.08: 119.5
| Std Dev. |  | 1.01: 2.7325; 1.02: 4.1673; 1.03: 1.0108; 1.04: 1.6723; 1.05: 2.9008; 1.06: 2.6077; 1.07: 1.3663; 1.08: 2.2509
| Rel.Std.Dev. |  | 1.01: 2.129%; 1.02: 3.661%; 1.03: 0.7709%; 1.04: 1.332%; 1.05: 2.13%; 1.06: 1.946%; 1.07: 1.081%; 1.08: 1.886%
| PDM3 |  | 1.01: 1.074%; 1.02: -10.35%; 1.03: 3.266%; 1.04: -1.134%; 1.05: 7.278%; 1.06: 5.537%; 1.07: -0.5011%; 1.08: -6.014%
| Z-Score (Absolute) |  | 1.01: 0.13; 1.02: 1.74; 1.03: 0.49; 1.04: 0.23; 1.05: 1.15; 1.06: 0.86; 1.07: 0.13; 1.08: 1.03
| Analytical results for Re in OREAS 627 (Indicative Value < 0.05 ppm) | [24 empty] | Indicative
Re | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab | 1
| No | No | 1.01: 1.15; 1.02: 1.17; 1.03: 1.20 | ppm
|  |  | 1.01: PF*MS; 1.02: PF*OES/MS; 1.03: PF*MS | 3
3
| 1 | 1 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 1
| 1 | 2 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 15
| 1 | 3 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 16
| 1 | 4 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | < 0.05
| 1 | 5 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02 | 21
| 1 | 6 | 1.01: < 0.1; 1.02: < 0.05; 1.03: < 0.02
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.03
| Analytical results for S in OREAS 627 (Certified Value 10.75 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 1.01 | 1.02 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.20 | 1.26 | 1.27 | [6 empty] | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | [6 empty] | 2
3
| 1 | 1 | 10.92 | 10.5 | 10.55 | 10.75 | 10.8 | 10.7 | 10.65 | 10.9 | 10.88 | 10.39 | 9.94 | 11.7958 | 8.49 | 10.1 | 11.2468 | 9.2 | 10.84 | [6 empty] | 1
| 1 | 2 | 10.8 | 10.6 | 10.7 | 10.7 | 10.95 | 10.6 | 10.8 | 11 | 10.69 | 10.56 | 10.03 | 11.3173 | 8.13 | 11 | 11.3307 | 8.95 | 11.32 | [6 empty] | 26
| 1 | 3 | 10.9 | 10.6 | 10.6 | 10.75 | 11 | 10.65 | 10.9 | 11.1 | 10.73 | 10.45 | 10.08 | 12.2831 | 8.28 | 11.2 | 10.9362 | 9.08 | 11.14 | [6 empty] | 16
| 1 | 4 | 10.97 | 10.5 | 10.65 | 10.75 | 11.25 | 10.75 | 10.65 | 11 | 10.66 | 10.79 | 9.96 | 12.3415 | 8.41 | 10.9 | 11.5304 | 8.97 | 10.86 | [6 empty] | 10.7512
| 1 | 5 | 10.86 | 10.5 | 10.65 | 10.75 | 11 | 10.55 | 10.6 | 11.1 | 10.78 | 10.15 | 9.86 | 11.2458 | 8.35 | 10.5 | 11.4727 | 9.05 | 10.96 | [6 empty] | 152
| 1 | 6 | 10.9 | 10.8 | 10.6 | 10.85 | 11 | 10.75 | 10.85 | 11 | 10.73 | 10.48 | 10.09 | 11.3229 | 8.36 | 10.8 | 11.0113 | 9.09 | 11
| Mean |  | 10.8917 | 10.5833 | 10.625 | 10.7583 | 11 | 10.6667 | 10.7417 | 11.0167 | 10.745 | 10.47 | 9.9933 | 11.7177 | 8.3367 | 10.75 | 11.2547 | 9.0567 | 11.02
| Median |  | 10.9 | 10.55 | 10.625 | 10.75 | 11 | 10.675 | 10.725 | 11 | 10.73 | 10.465 | 9.995 | 11.5594 | 8.355 | 10.85 | 11.2888 | 9.065 | 10.98
| Std Dev. |  | 0.057417 | 0.1169 | 0.05244 | 0.04916 | 0.14491 | 0.08165 | 0.12416 | 0.075277 | 0.077653 | 0.20967 | 0.089368 | 0.50075 | 0.12291 | 0.3937 | 0.24086 | 0.090701 | 0.18243
| Rel.Std.Dev. |  | 0.5272% | 1.105% | 0.4936% | 0.4569% | 1.317% | 0.7655% | 1.156% | 0.6833% | 0.7227% | 2.003% | 0.8943% | 4.273% | 1.474% | 3.662% | 2.14% | 1.001% | 1.655%
| PDM3 |  | 1.307% | -1.561% | -1.174% | 0.06665% | 2.314% | -0.786% | -0.08837% | 2.469% | -0.05737% | -2.615% | -7.049% | 8.99% | -22.46% | -0.01086% | 4.683% | -15.76% | 2.5%
| Z-Score (Absolute) |  | 0.39 | 0.43 | 0.32 | 0.04 | 0.67 | 0.21 | 0.01 | 0.72 | 0 | 0.73 | 1.99 | 2.57 | 6.37 | 0.01 | 1.35 | 4.46 | 0.73
| Analytical results for Sb in OREAS 627 (Certified Value 148 ppm) | [24 empty] | Certified
Sb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | [6 empty] | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.04: 1.09; 1.05: 1.10; 1.06: 1.15; 1.07: 1.16; 1.08: 1.23; 1.09: 1.25; 1.10: 1.26; 1.11: 1.28 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*OES; 1.10: PF*OES; 1.11: PF*OES/MS | [6 empty] | 0
0
| 1 | 1 | 1.01: 356; 1.02: 138; 1.04: 137; 1.05: 143.6; 1.06: 160.4418; 1.07: 155.6; 1.08: 143; 1.09: 107; 1.10: 150; 1.11: 100 | [6 empty] | 1
| 1 | 2 | 1.01: 356; 1.02: 139; 1.04: 146; 1.05: 143.6; 1.06: 163.6728; 1.07: 154.7; 1.08: 148; 1.09: 109; 1.10: 140; 1.11: 100 | [6 empty] | 45
| 1 | 3 | 1.01: 358; 1.02: 142; 1.04: 151; 1.05: 143.9; 1.06: 164.5224; 1.07: 153.6; 1.08: 147; 1.09: 106; 1.10: 150; 1.11: 100 | [6 empty] | 16
| 1 | 4 | 1.01: 358; 1.02: 136; 1.04: 150; 1.05: 146.4; 1.06: 164.274; 1.07: 155.3; 1.08: 149; 1.09: 106; 1.10: 140; 1.11: 200 | [6 empty] | 147.7445
| 1 | 5 | 1.01: 357; 1.02: 143; 1.04: 142; 1.05: 142; 1.06: 158.9113; 1.07: 158; 1.08: 156; 1.09: 109; 1.10: 140; 1.11: < 100 | [6 empty] | 153
| 1 | 6 | 1.01: 357; 1.02: 148; 1.04: 150; 1.05: 141.9; 1.06: 158.6121; 1.07: 157.7; 1.08: 150; 1.09: 109; 1.10: 150; 1.11: 200
| Mean |  | 1.01: 357; 1.02: 141; 1.04: 146; 1.05: 143.5667; 1.06: 161.739; 1.07: 155.8167; 1.08: 148.8333; 1.09: 107.6667; 1.10: 145; 1.11: 140
| Median |  | 1.01: 357; 1.02: 140.5; 1.04: 148; 1.05: 143.6; 1.06: 162.0573; 1.07: 155.45; 1.08: 148.5; 1.09: 108; 1.10: 145; 1.11: 100
| Std Dev. |  | 1.01: 0.89443; 1.02: 4.2895; 1.04: 5.5498; 1.05: 1.6354; 1.06: 2.7339; 1.07: 1.7198; 1.08: 4.2622; 1.09: 1.5055; 1.10: 5.4772; 1.11: 54.7723
| Rel.Std.Dev. |  | 1.01: 0.2505%; 1.02: 3.042%; 1.04: 3.801%; 1.05: 1.139%; 1.06: 1.69%; 1.07: 1.104%; 1.08: 2.864%; 1.09: 1.398%; 1.10: 3.777%; 1.11: 39.12%
| PDM3 |  | 1.01: 141.6%; 1.02: -4.565%; 1.04: -1.181%; 1.05: -2.828%; 1.06: 9.472%; 1.07: 5.464%; 1.08: 0.737%; 1.09: -27.13%; 1.10: -1.858%; 1.11: -5.242%
| Z-Score (Absolute) |  | 1.01: 19.25; 1.02: 0.41; 1.04: 0.05; 1.05: 0.18; 1.06: 1.48; 1.07: 0.94; 1.08: 0.3; 1.09: 3.44; 1.10: 0.05; 1.11: 1.87
| Analytical results for Sc in OREAS 627 (Certified Value 5.49 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 1.01 | 1.09 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | [12 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | [12 empty] | 2
2
| 1 | 1 | 261 | < 10 | < 10 | < 20 | 4.8 | 3 | 6.027 | 6 | < 5 | < 5 | 5 | [12 empty] | 1
| 1 | 2 | 259 | < 10 | < 10 | < 20 | 4.7 | 3 | 5.886 | 6 | < 5 | < 5 | 6 | [12 empty] | 46
| 1 | 3 | 261 | < 10 | < 10 | < 20 | 4.8 | 3 | 5.639 | 6 | 5 | < 5 | 6 | [12 empty] | 16
| 1 | 4 | 260 | < 10 | < 10 | < 20 | 4.7 | 3 | 5.983 | 7 | < 5 | < 5 | 6 | [12 empty] | 5.4947
| 1 | 5 | 259 | < 10 | < 10 | < 20 | 4.9 | 3 | 5.842 | 6 | < 5 | < 5 | 5 | [12 empty] | 154
| 1 | 6 | 260 | < 10 | < 10 | < 20 | 4.7 | 3 | 5.863 | 6 | < 5 | < 5 | 6
| Mean |  | 1.01: 260; 1.06: 4.7667; 1.07: 3; 1.08: 5.8733; 1.09: 6.1667; 1.10: 5; 1.12: 5.6667
| Median |  | 1.01: 260; 1.06: 4.75; 1.07: 3; 1.08: 5.8745; 1.09: 6; 1.10: 5; 1.12: 6
| Std Dev. |  | 1.01: 0.89443; 1.06: 0.08165; 1.07: 0; 1.08: 0.1355; 1.09: 0.40825; 1.12: 0.5164
| Rel.Std.Dev. |  | 1.01: 0.344%; 1.06: 1.713%; 1.07: 0%; 1.08: 2.307%; 1.09: 6.62%; 1.12: 9.113%
| PDM3 |  | 1.01: 4632%; 1.06: -13.25%; 1.07: -45.4%; 1.08: 6.892%; 1.09: 12.23%; 1.10: -9.003%; 1.12: 3.13%
| Z-Score (Absolute) |  | 147.38 | 0 | 0 | 2.89 | 0.13 | 1.16 | 0.5 | 0.67 | 1.2 | 1.44 | 0.39
| Analytical results for Se in OREAS 627 (Indicative Value < 5 ppm) | [24 empty] | Indicative
Se | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.15 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: < 8; 1.02: < 5; 1.04: < 20 | [6 empty] | 1
| 1 | 2 | 1.01: 8; 1.02: < 5; 1.04: < 20 | [6 empty] | 16
| 1 | 3 | 1.01: 11; 1.02: < 5; 1.04: < 20 | [6 empty] | 16
| 1 | 4 | 1.01: < 8; 1.02: < 5; 1.04: < 20 | [6 empty] | < 5
| 1 | 5 | 1.01: < 8; 1.02: < 5; 1.04: < 20 | [6 empty] | 22
| 1 | 6 | 1.01: < 8; 1.02: < 5; 1.04: < 20
| Mean |  | 1.01: 9.5
| Median |  | 1.01: 9.5
| Std Dev. |  | 1.01: 2.1213
| Rel.Std.Dev. |  | 1.01: 22.33%
PDM3
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.67; 1.04: 0.84
| Analytical results for Si in OREAS 627 (Certified Value 22.44 wt.%) | [24 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.12 | 1.15 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [7 empty] | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 22.2 | > 15 | 20.9 | 22.81 | 21.88 | 22.72 | 22.58 | 22.1 | 22.5 | 22.7 | 23.3703 | 22.5395 | 21.1 | 23.2 | 21.32 | 22.7 | [7 empty] | 1
| 1 | 2 | 22.4 | > 15 | 21.22 | 23.19 | 22.11 | 22.39 | 22.72 | 22.4 | 22.3 | 22 | 23.7351 | 22.4992 | 22 | 23.4 | 21.34 | 22.5 | [7 empty] | 0
| 1 | 3 | 22.5 | > 15 | 21.22 | 22.81 | 22.39 | 22.58 | 23.19 | 22.6 | 22.4 | 22.4 | 23.5242 | 21.3095 | 21.5 | 24 | 21.19 | 22.8 | [7 empty] | 16
| 1 | 4 | 22.2 | > 15 | 21.08 | 22.39 | 22.39 | 22.81 | 22.58 | 22.3 | 22.5 | 21.9 | 23.653 | 22.7751 | 22.2 | 23.3 | 21.09 | 22.3 | [7 empty] | 22.4397
| 1 | 5 | 22.2 | > 15 | 20.99 | 22.81 | 22.2 | 22.58 | 22.72 | 22.6 | 22.3 | 22.8 | 23.3814 | 22.9238 | 22.3 | 23.4 | 21.34 | 22.9 | [7 empty] | 155
| 1 | 6 | 22.6 | > 15 | 21.08 | 23.09 | 22.11 | 22.72 | 22.81 | 22.4 | 22.2 | 22.1 | 23.4921 | 22.0689 | 22.1 | 23.2 | 21.15 | 22.4
| Mean |  | 22.35 |  | 21.0817 | 22.85 | 22.18 | 22.6333 | 22.7667 | 22.4 | 22.3667 | 22.3167 | 23.526 | 22.3527 | 21.8667 | 23.4167 | 21.2383 | 22.6
| Median |  | 22.3 |  | 21.08 | 22.81 | 22.155 | 22.65 | 22.72 | 22.4 | 22.35 | 22.25 | 23.5081 | 22.5193 | 22.05 | 23.35 | 21.255 | 22.6
| Std Dev. |  | 0.17607 |  | 0.12624 | 0.27914 | 0.19411 | 0.14909 | 0.22589 | 0.18974 | 0.12111 | 0.37639 | 0.1457 | 0.58817 | 0.46762 | 0.29944 | 0.10907 | 0.23664
| Rel.Std.Dev. |  | 0.7878% |  | 0.5988% | 1.222% | 0.8752% | 0.6587% | 0.9922% | 0.847% | 0.5415% | 1.687% | 0.6193% | 2.631% | 2.138% | 1.279% | 0.5136% | 1.047%
| PDM3 |  | -0.3999% |  | -6.052% | 1.828% | -1.157% | 0.8628% | 1.457% | -0.177% | -0.3256% | -0.5484% | 4.841% | -0.388% | -2.554% | 4.354% | -5.354% | 0.7142%
| Z-Score (Absolute) |  | 0.04 | NA | 3.25 | 1.23 | 0.47 | 0.67 | 1.01 | 0.08 | 0 | 0.13 | 2.93 | 0.04 | 1.27 | 2.66 | 2.86 | 0.59
| Analytical results for Sm in OREAS 627 (Certified Value 5.72 ppm) | [24 empty] | Certified
Sm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.19; 1.07: 1.20; 1.08: 1.22; 1.09: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: BF*MS; 1.07: PF*MS; 1.08: PF*MS; 1.09: PF*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 5.5; 1.02: 5.2; 1.04: 6.1; 1.05: 5.7372; 1.06: 6.279; 1.07: 7; 1.08: 5.068; 1.09: 5.8 | [6 empty] | 1
| 1 | 2 | 1.01: 5.4; 1.02: 5.4; 1.04: 5.9; 1.05: 5.7646; 1.06: 6.317; 1.07: 5; 1.08: 5.4874; 1.09: 5.7 | [6 empty] | 8
| 1 | 3 | 1.01: 5.6; 1.02: 5.5; 1.04: 6; 1.05: 5.8437; 1.06: 6.346; 1.07: 6; 1.08: 5.4636; 1.09: 6 | [6 empty] | 16
| 1 | 4 | 1.01: 4.2; 1.02: 5.6; 1.04: 6.3; 1.05: 5.6477; 1.06: 6.322; 1.07: 5; 1.08: 5.2564; 1.09: 5.8 | [6 empty] | 5.7198
| 1 | 5 | 1.01: 4.9; 1.02: 5.4; 1.04: 6.1; 1.05: 5.8133; 1.06: 6.338; 1.07: 6; 1.08: 5.528; 1.09: 6 | [6 empty] | 156
| 1 | 6 | 1.01: 4.9; 1.02: 5.5; 1.04: 5.3; 1.05: 5.5892; 1.06: 6.31; 1.07: 5; 1.08: 5.2796; 1.09: 5.9
| Mean |  | 1.01: 5.0833; 1.02: 5.4333; 1.04: 5.95; 1.05: 5.7326; 1.06: 6.3187; 1.07: 5.6667; 1.08: 5.3472; 1.09: 5.8667
| Median |  | 1.01: 5.15; 1.02: 5.45; 1.04: 6.05; 1.05: 5.7509; 1.06: 6.3195; 1.07: 5.5; 1.08: 5.3716; 1.09: 5.85
| Std Dev. |  | 1.01: 0.52694; 1.02: 0.13663; 1.04: 0.34496; 1.05: 0.097641; 1.06: 0.023594; 1.07: 0.8165; 1.08: 0.17701; 1.09: 0.12111
| Rel.Std.Dev. |  | 1.01: 10.37%; 1.02: 2.515%; 1.04: 5.798%; 1.05: 1.703%; 1.06: 0.3734%; 1.07: 14.41%; 1.08: 3.31%; 1.09: 2.064%
| PDM3 |  | 1.01: -11.13%; 1.02: -5.008%; 1.04: 4.025%; 1.05: 0.2247%; 1.06: 10.47%; 1.07: -0.9286%; 1.08: -6.515%; 1.09: 2.568%
| Z-Score (Absolute) |  | 1.01: 1.46; 1.02: 0.67; 1.04: 0.49; 1.05: 0; 1.06: 1.32; 1.07: NA; 1.08: 0.87; 1.09: 0.3
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Sn in OREAS 627 (Indicative Value 6.94 ppm) | [24 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.02 | 1.03 | 1.09 | 1.10 | 1.11 | 1.13 | 1.15 | 1.16 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*MS | PF*OES | PF*OES | PF*OES | [10 empty] | 2
2
| 1 | 1 | 5.7 | 300 | 7 | 7 | < 50 | 4.0639 | < 100 | < 5 | NA | 5 | < 50 | < 50 | < 100 | [10 empty] | 1
| 1 | 2 | 19.6 | < 100 | 7 | 7 | < 50 | 4.495 | < 100 | < 5 | NA | 5 | < 50 | < 50 | < 100 | [10 empty] | 12
| 1 | 3 | 14.1 | < 100 | 7 | 7 | < 50 | 4.1987 | < 100 | < 5 | NA | 5 | < 50 | < 50 | < 100 | [10 empty] | 16
| 1 | 4 | 8.9 | < 100 | 7 | 7 | < 50 | 4.0982 | < 100 | < 5 | NA | 6 | < 50 | < 50 | < 100 | [10 empty] | 6.9382
| 1 | 5 | 13.1 | < 100 | 7 | 6 | < 50 | 3.4825 | < 100 | < 5 | NA | 6 | < 50 | < 50 | < 100 | [10 empty] | 18
| 1 | 6 | 16.2 | < 100 | 7 | 6 | < 50 | 3.2078 | < 100 | < 5 | NA | 6 | < 50 | < 50 | < 100
| Mean |  | 1.01: 12.9333; 1.02: 300; 1.04: 7; 1.05: 6.6667; 1.07: 3.9244; 1.11: 5.5
| Median |  | 1.01: 13.6; 1.02: 300; 1.04: 7; 1.05: 7; 1.07: 4.0811; 1.11: 5.5
| Std Dev. |  | 1.01: 5.0019; 1.04: 0; 1.05: 0.5164; 1.07: 0.48154; 1.11: 0.54772
| Rel.Std.Dev. |  | 1.01: 38.67%; 1.04: 0%; 1.05: 7.746%; 1.07: 12.27%; 1.11: 9.959%
| PDM3 |  | 1.01: 86.41%; 1.02: 4224%; 1.04: 0.8907%; 1.05: -3.914%; 1.07: -43.44%; 1.11: -20.73%
| Z-Score (Absolute) |  | 0.32 | 3.81 | 0.63 | 0.64 | 0.32 | 0.79 | 1.62 | 0.86 | NA | 0.7 | 0.32 | 0.32 | 1.62
| Analytical results for Sr in OREAS 627 (Certified Value 113 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.10 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.22 | 1.23 | 1.25 | 1.26 | [9 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | [9 empty] | 0
0
| 1 | 1 | 112 | 115 | 109 | 113 | 122 | 121.0315 | 118.2341 | 101 | 128 | 112 | 139.252 | 109 | 110 | 110 | [9 empty] | 1
| 1 | 2 | 115 | 113 | 117 | 114 | 123 | 124.6778 | 118.8017 | 103.6 | 127 | 107 | 134.609 | 109 | 111 | 110 | [9 empty] | 21
| 1 | 3 | 112 | 112 | 120 | 113 | 121 | 121.8964 | 119.5192 | 101.9 | 128 | 106 | 138.329 | 109 | 110 | 110 | [9 empty] | 16
| 1 | 4 | 112 | 119 | 112 | 112 | 123 | 123.7314 | 118.7883 | 103.2 | 126 | 110 | 139.885 | 108 | 110 | 110 | [9 empty] | 113.0918
| 1 | 5 | 112 | 111 | 119 | 114 | 121 | 124.142 | 117.2044 | 102.3 | 135 | 114 | 140.642 | 108 | 111 | 110 | [9 empty] | 157
| 1 | 6 | 112 | 112 | 111 | 114 | 121 | 121.0797 | 117.5018 | 102.4 | 130 | 111 | 141.32 | 108 | 109 | 110
| Mean |  | 112.5 | 113.6667 | 114.6667 | 113.3333 | 121.8333 | 122.7598 | 118.3416 | 102.4 | 129 | 110 | 139.0062 | 108.5 | 110.1667 | 110
| Median |  | 112 | 112.5 | 114.5 | 113.5 | 121.5 | 122.8139 | 118.5112 | 102.35 | 128 | 110.5 | 139.5685 | 108.5 | 110 | 110
| Std Dev. |  | 1.2247 | 2.9439 | 4.5898 | 0.8165 | 0.98319 | 1.6179 | 0.8727 | 0.92736 | 3.2249 | 3.0332 | 2.394 | 0.54772 | 0.75277 | 0
| Rel.Std.Dev. |  | 1.089% | 2.59% | 4.003% | 0.7204% | 0.807% | 1.318% | 0.7374% | 0.9056% | 2.5% | 2.757% | 1.722% | 0.5048% | 0.6833% | 0%
| PDM3 |  | -0.5233% | 0.5083% | 1.393% | 0.2136% | 7.73% | 8.549% | 4.642% | -9.454% | 14.07% | -2.734% | 22.91% | -4.06% | -2.586% | -2.734%
| Z-Score (Absolute) |  | 0.16 | 0.03 | 0.19 | 0.03 | 1.35 | 1.5 | 0.78 | 1.79 | 2.51 | 0.57 | 4.12 | 0.81 | 0.54 | 0.57
| Analytical results for Ta in OREAS 627 (Indicative Value 0.85 ppm) | [24 empty] | Indicative
Ta | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.20; 1.08: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: 1; 1.02: < 0.5; 1.04: 0.8; 1.05: 1.1668; 1.06: 1.6575; 1.07: NA; 1.08: 0.8 | [6 empty] | 1
| 1 | 2 | 1.01: 0.9; 1.02: < 0.5; 1.04: 0.6; 1.05: 1.119; 1.06: 1.6857; 1.07: NA; 1.08: 0.8 | [6 empty] | 13
| 1 | 3 | 1.01: 1.3; 1.02: < 0.5; 1.04: 0.6; 1.05: 1.1072; 1.06: 1.8205; 1.07: NA; 1.08: 0.8 | [6 empty] | 16
| 1 | 4 | 1.01: 0.9; 1.02: < 0.5; 1.04: 0.7; 1.05: 1.0421; 1.06: 1.865; 1.07: NA; 1.08: 0.8 | [6 empty] | 0.85322
| 1 | 5 | 1.01: 0.8; 1.02: < 0.5; 1.04: 0.5; 1.05: 1.0813; 1.06: 1.7932; 1.07: NA; 1.08: 0.8 | [6 empty] | 19
| 1 | 6 | 1.01: 0.8; 1.02: < 0.5; 1.04: 0.6; 1.05: 1.0811; 1.06: 1.6633; 1.07: NA; 1.08: 0.8
| Mean |  | 1.01: 0.95; 1.04: 0.63333; 1.05: 1.0996; 1.06: 1.7475; 1.08: 0.8
| Median |  | 1.01: 0.9; 1.04: 0.6; 1.05: 1.0942; 1.06: 1.7395; 1.08: 0.8
| Std Dev. |  | 1.01: 0.18708; 1.04: 0.10328; 1.05: 0.042264; 1.06: 0.089671; 1.08: 1.2162e-16
| Rel.Std.Dev. |  | 1.01: 19.69%; 1.04: 16.31%; 1.05: 3.844%; 1.06: 5.131%; 1.08: 1.52e-14%
| PDM3 |  | 1.01: 11.34%; 1.04: -25.77%; 1.05: 28.87%; 1.06: 104.8%; 1.08: -6.238%
| Z-Score (Absolute) |  | 1.01: 0.22; 1.02: 1.81; 1.04: 0.7; 1.05: 0.65; 1.06: 2.52; 1.07: NA; 1.08: 0.22
| Analytical results for Tb in OREAS 627 (Certified Value 0.68 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.19; 1.07: 1.20; 1.08: 1.22; 1.09: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: BF*MS; 1.07: PF*MS; 1.08: PF*MS; 1.09: PF*MS | [6 empty] | 2
3
| 1 | 1 | 1.01: 0.8; 1.02: 0.58; 1.04: 0.7; 1.05: 0.72039; 1.06: 0.725; 1.07: < 2; 1.08: 0.7458; 1.09: 0.69 | [6 empty] | 1
| 1 | 2 | 1.01: 0.7; 1.02: 0.62; 1.04: 0.7; 1.05: 0.68931; 1.06: 0.728; 1.07: < 2; 1.08: 0.6567; 1.09: 0.67 | [6 empty] | 36
| 1 | 3 | 1.01: 0.7; 1.02: 0.61; 1.04: 0.6; 1.05: 0.67911; 1.06: 0.726; 1.07: < 2; 1.08: 0.59173; 1.09: 0.74 | [6 empty] | 16
| 1 | 4 | 1.01: 0.6; 1.02: 0.65; 1.04: 0.6; 1.05: 0.67945; 1.06: 0.727; 1.07: < 2; 1.08: 0.73524; 1.09: 0.7 | [6 empty] | 0.68256
| 1 | 5 | 1.01: 0.6; 1.02: 0.59; 1.04: 0.6; 1.05: 0.69093; 1.06: 0.731; 1.07: < 2; 1.08: 0.65079; 1.09: 0.73 | [6 empty] | 158
| 1 | 6 | 1.01: 0.6; 1.02: 0.64; 1.04: 0.6; 1.05: 0.73799; 1.06: 0.724; 1.07: < 2; 1.08: 0.61835; 1.09: 0.7
| Mean |  | 1.01: 0.66667; 1.02: 0.615; 1.04: 0.63333; 1.05: 0.69953; 1.06: 0.72683; 1.08: 0.66643; 1.09: 0.705
| Median |  | 1.01: 0.65; 1.02: 0.615; 1.04: 0.6; 1.05: 0.69012; 1.06: 0.7265; 1.08: 0.65374; 1.09: 0.7
| Std Dev. |  | 1.01: 0.08165; 1.02: 0.027386; 1.04: 0.05164; 1.05: 0.024138; 1.06: 0.0024833; 1.08: 0.062074; 1.09: 0.025884
| Rel.Std.Dev. |  | 1.01: 12.25%; 1.02: 4.453%; 1.04: 8.154%; 1.05: 3.451%; 1.06: 0.3417%; 1.08: 9.314%; 1.09: 3.672%
| PDM3 |  | 1.01: -2.328%; 1.02: -9.898%; 1.04: -7.212%; 1.05: 2.487%; 1.06: 6.486%; 1.08: -2.362%; 1.09: 3.288%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 1.95; 1.04: NA; 1.05: 0.06; 1.06: 0.55; 1.07: 6.65; 1.08: 0.8; 1.09: 0.06
Results from laboratories 1.02 and 1.10 were removed due to their 0.1 ppm reading resolution.
| Analytical results for Te in OREAS 627 (Indicative Value < 5 ppm) | [24 empty] | Indicative
Te | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.15 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS | [6 empty] | 2
2
| 1 | 1 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 1
| 1 | 2 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 14
| 1 | 3 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 16
| 1 | 4 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | < 5
| 1 | 5 | 1.01: < 6; 1.02: < 5; 1.04: < 1 | [6 empty] | 20
| 1 | 6 | 1.01: < 6; 1.02: < 5; 1.04: < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 1.01 to 1.04
| Analytical results for Th in OREAS 627 (Certified Value 12.1 ppm) | [24 empty] | Certified
Th | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab | [6 empty] | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.16; 1.08: 1.17; 1.09: 1.19; 1.10: 1.20; 1.11: 1.23 | [6 empty] | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*OES/MS; 1.09: BF*MS; 1.10: PF*MS; 1.11: PF*MS | [6 empty] | 1
2
| 1 | 1 | 1.01: 11.1; 1.02: 10.2; 1.04: 12.6; 1.05: 11.982; 1.06: 13.8803; 1.07: 12.1; 1.08: 14.1; 1.09: 12.04; 1.10: 11; 1.11: 11.4 | [6 empty] | 1
| 1 | 2 | 1.01: 10.4; 1.02: 10.4; 1.04: 12.6; 1.05: 11.47; 1.06: 13.4581; 1.07: 11.9; 1.08: 13.3; 1.09: 12.215; 1.10: 11; 1.11: 11.2 | [6 empty] | 38
| 1 | 3 | 1.01: 10.6; 1.02: 10.8; 1.04: 12.6; 1.05: 11.6247; 1.06: 14.0951; 1.07: 12.1; 1.08: 13.3; 1.09: 12.195; 1.10: 11; 1.11: 12 | [6 empty] | 16
| 1 | 4 | 1.01: 10.2; 1.02: 11.1; 1.04: 12.5; 1.05: 11.4852; 1.06: 13.5362; 1.07: 12.1; 1.08: 13.6; 1.09: 12.224; 1.10: 11; 1.11: 11.7 | [6 empty] | 12.0613
| 1 | 5 | 1.01: 10.7; 1.02: 10.3; 1.04: 12.5; 1.05: 12.004; 1.06: 13.2423; 1.07: 11.6; 1.08: 14.2; 1.09: 12.265; 1.10: 11; 1.11: 12.3 | [6 empty] | 159
| 1 | 6 | 1.01: 10.5; 1.02: 11; 1.04: 12.3; 1.05: 11.6191; 1.06: 12.9421; 1.07: 11.9; 1.08: 13.5; 1.09: 12.234; 1.10: 11; 1.11: 12.1
| Mean |  | 1.01: 10.5833; 1.02: 10.6333; 1.04: 12.5167; 1.05: 11.6975; 1.06: 13.5257; 1.07: 11.95; 1.08: 13.6667; 1.09: 12.1955; 1.10: 11; 1.11: 11.7833
| Median |  | 1.01: 10.55; 1.02: 10.6; 1.04: 12.55; 1.05: 11.6219; 1.06: 13.4972; 1.07: 12; 1.08: 13.55; 1.09: 12.2195; 1.10: 11; 1.11: 11.85
| Std Dev. |  | 1.01: 0.30605; 1.02: 0.38297; 1.04: 0.1169; 1.05: 0.23796; 1.06: 0.41847; 1.07: 0.19748; 1.08: 0.39328; 1.09: 0.079606; 1.10: 0; 1.11: 0.42622
| Rel.Std.Dev. |  | 1.01: 2.892%; 1.02: 3.602%; 1.04: 0.934%; 1.05: 2.034%; 1.06: 3.094%; 1.07: 1.653%; 1.08: 2.878%; 1.09: 0.6527%; 1.10: 0%; 1.11: 3.617%
| PDM3 |  | 1.01: -12.25%; 1.02: -11.84%; 1.04: 3.775%; 1.05: -3.017%; 1.06: 12.14%; 1.07: -0.9231%; 1.08: 13.31%; 1.09: 1.112%; 1.10: -8.799%; 1.11: -2.305%
| Z-Score (Absolute) |  | 1.01: 1.63; 1.02: 1.57; 1.04: 0.67; 1.05: 0.3; 1.06: 1.87; 1.07: 0; 1.08: 2.04; 1.09: 0.29; 1.10: NA; 1.11: 0.2
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for Ti in OREAS 627 (Certified Value 0.152 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 |  |  | wt.%
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES |  |  | 3
3
| 1 | 1 | 0.16 | 0.15 | 0.16 | 0.15 | 0.156 | 0.15 | 0.15 | 0.15 | 0.15 | 0.16 | 0.15 | 0.16 | 0.15932 | 0.16094 | 0.14 | 0.154 | 0.1575 | 0.15 | 0.16 | 0.15 | 0.16 |  |  | 1
| 1 | 2 | 0.16 | 0.15 | 0.16 | 0.15 | 0.156 | 0.15 | 0.15 | 0.156 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15796 | 0.15888 | 0.14 | 0.149 | 0.1559 | 0.15 | 0.16 | 0.15 | 0.16 |  |  | 39
| 1 | 3 | 0.16 | 0.15 | 0.16 | 0.15 | 0.156 | 0.15 | 0.15 | 0.156 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15842 | 0.15258 | 0.16 | 0.149 | 0.1503 | 0.15 | 0.16 | 0.15 | 0.15 |  |  | 16
| 1 | 4 | 0.17 | 0.15 | 0.16 | 0.15 | 0.15 | 0.156 | 0.15 | 0.156 | 0.15 | 0.15 | 0.16 | 0.15 | 0.15824 | 0.14479 | 0.14 | 0.15 | 0.1599 | 0.15 | 0.16 | 0.15 | 0.15 |  |  | 0.15157
| 1 | 5 | 0.16 | 0.15 | 0.16 | 0.15 | 0.15 | 0.156 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15 | 0.16136 | 0.15176 | 0.16 | 0.15 | 0.1582 | 0.15 | 0.16 | 0.15 | 0.16 |  |  | 160
| 1 | 6 | 0.16 | 0.16 | 0.16 | 0.15 | 0.156 | 0.15 | 0.15 | 0.174 | 0.16 | 0.15 | 0.15 | 0.15 | 0.15638 | 0.14202 | 0.14 | 0.15 | 0.1518 | 0.15 | 0.16 | 0.15 | 0.15
| Mean |  | 0.16167 | 0.15167 | 0.16 | 0.15 | 0.154 | 0.152 | 0.15 | 0.157 | 0.15167 | 0.15167 | 0.15167 | 0.15167 | 0.15861 | 0.15183 | 0.14667 | 0.15033 | 0.1556 | 0.15 | 0.16 | 0.15 | 0.155
| Median |  | 0.16 | 0.15 | 0.16 | 0.15 | 0.156 | 0.15 | 0.15 | 0.156 | 0.15 | 0.15 | 0.15 | 0.15 | 0.15833 | 0.15217 | 0.14 | 0.15 | 0.1567 | 0.15 | 0.16 | 0.15 | 0.155
| Std Dev. |  | 0.0040825 | 0.0040825 | 0 | 0 | 0.0030984 | 0.0030984 | 0 | 0.0088318 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0040825 | 0.0016499 | 0.007472 | 0.010328 | 0.0018619 | 0.003781 | 0 | 0 | 0 | 0.0054772
| Rel.Std.Dev. |  | 2.525% | 2.692% | 0% | 0% | 2.012% | 2.038% | 0% | 5.625% | 2.692% | 2.692% | 2.692% | 2.692% | 1.04% | 4.921% | 7.042% | 1.239% | 2.43% | 0% | 0% | 0% | 3.534%
| PDM3 |  | 6.661% | 0.06301% | 5.561% | -1.037% | 1.602% | 0.2829% | -1.037% | 3.582% | 0.06301% | 0.06301% | 0.06301% | 0.06301% | 4.646% | 0.1702% | -3.236% | -0.8167% | 2.658% | -1.037% | 5.561% | -1.037% | 2.262%
| Z-Score (Absolute) |  | 3.76 | 0 | 3.13 | 0.67 | 0.83 | 0.08 | 0.67 | 1.95 | 0 | 0 | 0 | 0 | 2.61 | 0.06 | 1.88 | 0.5 | 1.48 | 0.63 | 3.13 | 0.63 | 1.25
| Analytical results for Tl in OREAS 627 (Certified Value 14.4 ppm) | [24 empty] | Certified
Tl | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.15; 1.05: 1.16; 1.06: 1.17; 1.07: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*MS; 1.05: PF*OES; 1.06: PF*OES/MS; 1.07: PF*MS |  |  | 1
2
| 1 | 1 | 1.01: 14.3; 1.02: 12.8; 1.03: 14.8; 1.04: 14.5178; 1.05: 15.5; 1.06: 16.1; 1.07: 13.1 |  |  | 1
| 1 | 2 | 1.01: 14.2; 1.02: 13.5; 1.03: 14.8; 1.04: 14.5422; 1.05: 15.4; 1.06: 15.6; 1.07: 12.8 |  |  | 0
| 1 | 3 | 1.01: 15.5; 1.02: 13.4; 1.03: 15.3; 1.04: 14.2233; 1.05: 15.6; 1.06: 15.8; 1.07: 13 |  |  | 16
| 1 | 4 | 1.01: 14.4; 1.02: 13.6; 1.03: 14.9; 1.04: 14.2206; 1.05: 15.7; 1.06: 15.3; 1.07: 13 |  |  | 14.4411
| 1 | 5 | 1.01: 14.2; 1.02: 13; 1.03: 14.6; 1.04: 14.2405; 1.05: 15.3; 1.06: 15.5; 1.07: 13.4 |  |  | 161
| 1 | 6 | 1.01: 14.4; 1.02: 13.7; 1.03: 14.6; 1.04: 14.2799; 1.05: 15.2; 1.06: 15.7; 1.07: 13.7
| Mean |  | 1.01: 14.5; 1.02: 13.3333; 1.03: 14.8333; 1.04: 14.3374; 1.05: 15.45; 1.06: 15.6667; 1.07: 13.1667
| Median |  | 1.01: 14.35; 1.02: 13.45; 1.03: 14.8; 1.04: 14.2602; 1.05: 15.45; 1.06: 15.65; 1.07: 13.05
| Std Dev. |  | 1.01: 0.498; 1.02: 0.3559; 1.03: 0.2582; 1.04: 0.15087; 1.05: 0.18708; 1.06: 0.27325; 1.07: 0.3266
| Rel.Std.Dev. |  | 1.01: 3.434%; 1.02: 2.669%; 1.03: 1.741%; 1.04: 1.052%; 1.05: 1.211%; 1.06: 1.744%; 1.07: 2.48%
| PDM3 |  | 1.01: 0.4082%; 1.02: -7.671%; 1.03: 2.716%; 1.04: -0.7179%; 1.05: 6.987%; 1.06: 8.487%; 1.07: -8.825%
| Z-Score (Absolute) |  | 1.01: 0; 1.02: 0.83; 1.03: 0.24; 1.04: 0.12; 1.05: 0.67; 1.06: 0.83; 1.07: 0.95
| Analytical results for Tm in OREAS 627 (Certified Value 0.23 ppm) | [24 empty] | Certified
Tm | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS |  |  | 2
2
| 1 | 1 | 1.01: 0.2; 1.02: 0.21; 1.03: 0.2; 1.04: 0.2511; 1.05: 0.234; 1.06: < 2; 1.07: 0.23 |  |  | 1
| 1 | 2 | 1.01: 0.3; 1.02: 0.21; 1.03: 0.3; 1.04: 0.23574; 1.05: 0.235; 1.06: < 2; 1.07: 0.23 |  |  | 9
| 1 | 3 | 1.01: 0.2; 1.02: 0.22; 1.03: 0.3; 1.04: 0.20293; 1.05: 0.244; 1.06: < 2; 1.07: 0.24 |  |  | 16
| 1 | 4 | 1.01: 0.2; 1.02: 0.24; 1.03: 0.3; 1.04: 0.25357; 1.05: 0.238; 1.06: < 2; 1.07: 0.23 |  |  | 0.23297
| 1 | 5 | 1.01: 0.2; 1.02: 0.22; 1.03: 0.2; 1.04: 0.26922; 1.05: 0.241; 1.06: < 2; 1.07: 0.24 |  |  | 162
| 1 | 6 | 1.01: 0.2; 1.02: 0.23; 1.03: 0.2; 1.04: 0.21731; 1.05: 0.245; 1.06: < 2; 1.07: 0.22
| Mean |  | 1.01: 0.21667; 1.02: 0.22167; 1.03: 0.25; 1.04: 0.23831; 1.05: 0.2395; 1.07: 0.23167
| Median |  | 1.01: 0.2; 1.02: 0.22; 1.03: 0.25; 1.04: 0.24342; 1.05: 0.2395; 1.07: 0.23
| Std Dev. |  | 1.01: 0.040825; 1.02: 0.01169; 1.03: 0.054772; 1.04: 0.024704; 1.05: 0.0045935; 1.07: 0.0075277
| Rel.Std.Dev. |  | 1.01: 18.84%; 1.02: 5.274%; 1.03: 21.91%; 1.04: 10.37%; 1.05: 1.918%; 1.07: 3.249%
| PDM3 |  | 1.01: -6.997%; 1.02: -4.851%; 1.03: 7.311%; 1.04: 2.293%; 1.05: 2.804%; 1.07: -0.5588%
| Z-Score (Absolute) |  | 1.01: 1.25; 1.02: 0.96; 1.03: 0.67; 1.04: 0; 1.05: 0.07; 1.06: 43.94; 1.07: 0.38
| Analytical results for U in OREAS 627 (Certified Value 5.23 ppm) | [24 empty] | Certified
U | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.15; 1.06: 1.17; 1.07: 1.19; 1.08: 1.20; 1.09: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*OES/MS; 1.07: BF*MS; 1.08: PF*MS; 1.09: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 4.9; 1.02: 4.4; 1.03: 5.1; 1.04: 5.3049; 1.05: 5.6108; 1.06: 5.98; 1.07: 5.226; 1.08: 6; 1.09: 5.2 |  |  | 1
| 1 | 2 | 1.01: 4.7; 1.02: 4.5; 1.03: 5; 1.04: 4.7679; 1.05: 5.5912; 1.06: 5.71; 1.07: 5.246; 1.08: 6; 1.09: 5.17 |  |  | 41
| 1 | 3 | 1.01: 4.7; 1.02: 4.8; 1.03: 5.2; 1.04: 5.8579; 1.05: 5.6161; 1.06: 5.5; 1.07: 5.215; 1.08: 6; 1.09: 5.3 |  |  | 16
| 1 | 4 | 1.01: 4.4; 1.02: 4.9; 1.03: 5.2; 1.04: 4.9464; 1.05: 5.7217; 1.06: 5.68; 1.07: 5.243; 1.08: 6; 1.09: 5.17 |  |  | 5.2325
| 1 | 5 | 1.01: 4.8; 1.02: 4.5; 1.03: 5.2; 1.04: 5.6629; 1.05: 5.7157; 1.06: 5.87; 1.07: 5.243; 1.08: 6; 1.09: 5.33 |  |  | 163
| 1 | 6 | 1.01: 4.7; 1.02: 4.7; 1.03: 4.5; 1.04: 6.4041; 1.05: 5.6444; 1.06: 5.61; 1.07: 5.245; 1.08: 6; 1.09: 5.54
| Mean |  | 1.01: 4.7; 1.02: 4.6333; 1.03: 5.0333; 1.04: 5.4907; 1.05: 5.65; 1.06: 5.725; 1.07: 5.2363; 1.08: 6; 1.09: 5.285
| Median |  | 1.01: 4.7; 1.02: 4.6; 1.03: 5.15; 1.04: 5.4839; 1.05: 5.6302; 1.06: 5.695; 1.07: 5.243; 1.08: 6; 1.09: 5.25
| Std Dev. |  | 1.01: 0.16733; 1.02: 0.19664; 1.03: 0.27325; 1.04: 0.60859; 1.05: 0.055922; 1.06: 0.17444; 1.07: 0.012801; 1.08: 0; 1.09: 0.14209
| Rel.Std.Dev. |  | 1.01: 3.56%; 1.02: 4.244%; 1.03: 5.429%; 1.04: 11.08%; 1.05: 0.9898%; 1.06: 3.047%; 1.07: 0.2445%; 1.08: 0%; 1.09: 2.689%
| PDM3 |  | 1.01: -10.18%; 1.02: -11.45%; 1.03: -3.807%; 1.04: 4.934%; 1.05: 7.978%; 1.06: 9.411%; 1.07: 0.07245%; 1.08: 14.67%; 1.09: 1.003%
| Z-Score (Absolute) |  | 1.01: 1.22; 1.02: 1.37; 1.03: 0.5; 1.04: 0.5; 1.05: 0.85; 1.06: 1.01; 1.07: 0.05; 1.08: NA; 1.09: 0.05
Results from laboratory 1.20 were removed due to their 1 ppm reading resolution.
| Analytical results for V in OREAS 627 (Certified Value 23 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.09 | 1.12 | 1.13 | 1.15 | 1.16 | 1.17 | 1.20 | 1.23 | 1.25 | 1.26 | [10 empty] | ppm
|  |  | PF*OES/MS | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES | PF*OES | [10 empty] | 1
2
| 1 | 1 | 25 | 23 | < 100 | 25 | 20 | 23.92 | 25.5041 | 30.6 | 30 | NA | 21 | 23 | 30 | [10 empty] | 1
| 1 | 2 | 26 | 23 | < 100 | 25 | < 20 | 24.6571 | 24.7395 | 30.1 | 30 | NA | 23 | 21 | 20 | [10 empty] | 23
| 1 | 3 | 25 | 20 | < 100 | 25 | 22 | 23.986 | 22.106 | 32 | 30 | NA | 24 | 21 | 20 | [10 empty] | 16
| 1 | 4 | 25 | 20 | < 100 | 25 | 22 | 24.7573 | 23.66 | 31.9 | 30 | NA | 21 | 20 | 30 | [10 empty] | 23.0373
| 1 | 5 | 25 | 22 | < 100 | 25 | 20 | 25.4899 | 24.229 | 33.5 | 30 | NA | 22 | 21 | 20 | [10 empty] | 164
| 1 | 6 | 25 | 20 | < 100 | 26 | 23 | 23.8824 | 22.4603 | 31.5 | 30 | NA | 21 | 20 | 20
| Mean |  | 1.01: 25.1667; 1.02: 21.3333; 1.04: 25.1667; 1.05: 21.4; 1.06: 24.4488; 1.07: 23.7831; 1.08: 31.6; 1.09: 30; 1.11: 22; 1.12: 21; 1.13: 23.3333
| Median |  | 1.01: 25; 1.02: 21; 1.04: 25; 1.05: 22; 1.06: 24.3215; 1.07: 23.9445; 1.08: 31.7; 1.09: 30; 1.11: 21.5; 1.12: 21; 1.13: 20
| Std Dev. |  | 1.01: 0.40825; 1.02: 1.5055; 1.04: 0.40825; 1.05: 1.3416; 1.06: 0.63828; 1.07: 1.3155; 1.08: 1.1933; 1.09: 0; 1.11: 1.2649; 1.12: 1.0954; 1.13: 5.164
| Rel.Std.Dev. |  | 1.01: 1.622%; 1.02: 7.057%; 1.04: 1.622%; 1.05: 6.269%; 1.06: 2.611%; 1.07: 5.531%; 1.08: 3.776%; 1.09: 0%; 1.11: 5.75%; 1.12: 5.216%; 1.13: 22.13%
| PDM3 |  | 1.01: 9.243%; 1.02: -7.397%; 1.04: 9.243%; 1.05: -7.107%; 1.06: 6.127%; 1.07: 3.237%; 1.08: 37.17%; 1.09: 30.22%; 1.11: -4.503%; 1.12: -8.844%; 1.13: 1.285%
| Z-Score (Absolute) |  | 0.29 | 0.77 | 7.13 | 0.29 | 1.27 | 0.09 | 0.09 | 2.06 | NA | NA | 0.58 | 0.86 | NA
| Analytical results for W in OREAS 627 (Certified Value 8.45 ppm) | [24 empty] | Certified
W | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.03; 1.04: 1.09; 1.05: 1.10; 1.06: 1.15; 1.07: 1.16; 1.08: 1.20; 1.09: 1.23; 1.10: 1.25; 1.11: 1.26 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*OES/MS; 1.05: PF*MS; 1.06: PF*MS; 1.07: PF*OES; 1.08: PF*MS; 1.09: PF*MS; 1.10: PF*OES; 1.11: PF*OES |  |  | 2
2
| 1 | 1 | 1.01: 1134; 1.02: 9.1; 1.03: < 100; 1.04: 7; 1.05: 9; 1.06: 9.7902; 1.07: 9.1; 1.08: 6; 1.09: 9; 1.10: 142; 1.11: < 50 |  |  | 1
| 1 | 2 | 1.01: 1150; 1.02: 8.3; 1.03: < 100; 1.04: 8; 1.05: 10; 1.06: 9.7531; 1.07: 9.2; 1.08: 6; 1.09: 9; 1.10: 144; 1.11: < 50 |  |  | 43
| 1 | 3 | 1.01: 1144; 1.02: 8.8; 1.03: < 100; 1.04: 8; 1.05: 9; 1.06: 9.8593; 1.07: 9.3; 1.08: 6; 1.09: 9; 1.10: 148; 1.11: < 50 |  |  | 16
| 1 | 4 | 1.01: 1112; 1.02: 8.7; 1.03: < 100; 1.04: 8; 1.05: 9; 1.06: 9.848; 1.07: 9.1; 1.08: 7; 1.09: 9; 1.10: 148; 1.11: < 50 |  |  | 8.4497
| 1 | 5 | 1.01: 1133; 1.02: 8.1; 1.03: < 100; 1.04: 8; 1.05: 7; 1.06: 9.7028; 1.07: 9.2; 1.08: 6; 1.09: 9; 1.10: 144; 1.11: < 50 |  |  | 165
| 1 | 6 | 1.01: 1128; 1.02: 8.1; 1.03: < 100; 1.04: 8; 1.05: 8; 1.06: 9.7357; 1.07: 9; 1.08: 9; 1.09: 9; 1.10: 143; 1.11: < 50
| Mean |  | 1.01: 1133.5; 1.02: 8.5167; 1.04: 7.8333; 1.05: 8.6667; 1.06: 9.7815; 1.07: 9.15; 1.08: 6.6667; 1.09: 9; 1.10: 144.8333
| Median |  | 1.01: 1133.5; 1.02: 8.5; 1.04: 8; 1.05: 9; 1.06: 9.7716; 1.07: 9.15; 1.08: 6; 1.09: 9; 1.10: 144
| Std Dev. |  | 1.01: 13.2325; 1.02: 0.41191; 1.04: 0.40825; 1.05: 1.0328; 1.06: 0.062687; 1.07: 0.10488; 1.08: 1.2111; 1.09: 0; 1.10: 2.5626
| Rel.Std.Dev. |  | 1.01: 1.167%; 1.02: 4.836%; 1.04: 5.212%; 1.05: 11.92%; 1.06: 0.6409%; 1.07: 1.146%; 1.08: 18.17%; 1.09: 0%; 1.10: 1.769%
| PDM3 |  | 1.01: 13315%; 1.02: 0.7921%; 1.04: -7.295%; 1.05: 2.567%; 1.06: 15.76%; 1.07: 8.287%; 1.08: -21.1%; 1.09: 6.512%; 1.10: 1614%
| Z-Score (Absolute) |  | 1.01: 575.82; 1.02: 0.32; 1.03: 20.92; 1.04: 0.67; 1.05: 0.25; 1.06: 0.32; 1.07: 0; 1.08: 1.27; 1.09: 0.08; 1.10: 69.49; 1.11: 8.12
| Analytical results for Y in OREAS 627 (Certified Value 18 ppm) | [24 empty] | Certified
Y | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab; 1.09: Lab; 1.10: Lab; 1.11: Lab; 1.12: Lab |  |  | 1
| No | No | 1.01: 1.01; 1.02: 1.02; 1.03: 1.09; 1.04: 1.10; 1.05: 1.13; 1.06: 1.15; 1.07: 1.17; 1.08: 1.19; 1.09: 1.20; 1.10: 1.23; 1.11: 1.25; 1.12: 1.26 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*OES/MS; 1.04: PF*MS; 1.05: PF*OES/MS; 1.06: PF*MS; 1.07: PF*OES/MS; 1.08: BF*MS; 1.09: PF*MS; 1.10: PF*MS; 1.11: PF*OES; 1.12: PF*OES |  |  | 1
2
| 1 | 1 | 1.01: 19; 1.02: 18.4; 1.03: 15.9; 1.04: 18; 1.05: 18.3898; 1.06: 19.261; 1.07: 19.6; 1.08: 17.286; 1.09: 32; 1.10: 17.5; 1.11: 17; 1.12: 17 |  |  | 1
| 1 | 2 | 1.01: 20; 1.02: 17.5; 1.03: 16.9; 1.04: 17.7; 1.05: 18.0157; 1.06: 19.3632; 1.07: 19.7; 1.08: 17.154; 1.09: 19; 1.10: 17.5; 1.11: 17; 1.12: 18 |  |  | 44
| 1 | 3 | 1.01: 20; 1.02: 18.3; 1.03: 17.4; 1.04: 17.8; 1.05: 17.9673; 1.06: 19.7526; 1.07: 20.3; 1.08: 17.101; 1.09: 18; 1.10: 17.8; 1.11: 18; 1.12: 17 |  |  | 16
| 1 | 4 | 1.01: 19; 1.02: 14.7; 1.03: 17.4; 1.04: 17.9; 1.05: 18.3756; 1.06: 19.7814; 1.07: 19.7; 1.08: 17.126; 1.09: 20; 1.10: 17.7; 1.11: 17; 1.12: 17 |  |  | 18.0326
| 1 | 5 | 1.01: 19; 1.02: 17.8; 1.03: 16.8; 1.04: 17.6; 1.05: 18.7527; 1.06: 19.0831; 1.07: 19.7; 1.08: 17.214; 1.09: 24; 1.10: 18.1; 1.11: 18; 1.12: 17 |  |  | 166
| 1 | 6 | 1.01: 20; 1.02: 16.2; 1.03: 17.3; 1.04: 17.3; 1.05: 18.5999; 1.06: 19.2854; 1.07: 19.2; 1.08: 17.256; 1.09: 18; 1.10: 18.1; 1.11: 17; 1.12: 17
| Mean |  | 1.01: 19.5; 1.02: 17.15; 1.03: 16.95; 1.04: 17.7167; 1.05: 18.3502; 1.06: 19.4211; 1.07: 19.7; 1.08: 17.1895; 1.09: 21.8333; 1.10: 17.7833; 1.11: 17.3333; 1.12: 17.1667
| Median |  | 1.01: 19.5; 1.02: 17.65; 1.03: 17.1; 1.04: 17.75; 1.05: 18.3827; 1.06: 19.3243; 1.07: 19.7; 1.08: 17.184; 1.09: 19.5; 1.10: 17.75; 1.11: 17; 1.12: 17
| Std Dev. |  | 1.01: 0.54772; 1.02: 1.4377; 1.03: 0.57533; 1.04: 0.24833; 1.05: 0.31147; 1.06: 0.28333; 1.07: 0.35214; 1.08: 0.074107; 1.09: 5.4559; 1.10: 0.27142; 1.11: 0.5164; 1.12: 0.40825
| Rel.Std.Dev. |  | 1.01: 2.809%; 1.02: 8.383%; 1.03: 3.394%; 1.04: 1.402%; 1.05: 1.697%; 1.06: 1.459%; 1.07: 1.787%; 1.08: 0.4311%; 1.09: 24.99%; 1.10: 1.526%; 1.11: 2.979%; 1.12: 2.378%
| PDM3 |  | 1.01: 8.137%; 1.02: -4.894%; 1.03: -6.004%; 1.04: -1.752%; 1.05: 1.761%; 1.06: 7.7%; 1.07: 9.247%; 1.08: -4.675%; 1.09: 21.08%; 1.10: -1.382%; 1.11: -3.878%; 1.12: -4.802%
| Z-Score (Absolute) |  | 1.01: NA; 1.02: 0.67; 1.03: 0.9; 1.04: 0.04; 1.05: 0.67; 1.06: 1.88; 1.07: 2.19; 1.08: 0.63; 1.09: NA; 1.10: 0.04; 1.11: NA; 1.12: NA
Results from laboratories 1.01, 1.20, 1.25 and 1.26 were removed due to their 1 ppm reading resolution.
| Analytical results for Yb in OREAS 627 (Certified Value 1.5 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab; 1.06: Lab; 1.07: Lab; 1.08: Lab |  |  | 1
| No | No | 1.01: 1.02; 1.02: 1.09; 1.03: 1.10; 1.04: 1.13; 1.05: 1.19; 1.06: 1.20; 1.07: 1.22; 1.08: 1.23 |  |  | ppm
|  |  | 1.01: PF*OES/MS; 1.02: PF*OES/MS; 1.03: PF*MS; 1.04: PF*OES/MS; 1.05: BF*MS; 1.06: PF*MS; 1.07: PF*MS; 1.08: PF*MS |  |  | 2
3
| 1 | 1 | 1.01: 1.3; 1.02: 1.4; 1.03: 1.6; 1.04: 1.5497; 1.05: 1.605; 1.06: < 2; 1.07: 1.5105; 1.08: 1.4 |  |  | 1
| 1 | 2 | 1.01: 1.5; 1.02: 1.5; 1.03: 1.5; 1.04: 1.5004; 1.05: 1.615; 1.06: < 2; 1.07: 1.358; 1.08: 1.5 |  |  | 45
| 1 | 3 | 1.01: 1.4; 1.02: 1.5; 1.03: 1.7; 1.04: 1.5054; 1.05: 1.629; 1.06: < 2; 1.07: 1.2401; 1.08: 1.5 |  |  | 16
| 1 | 4 | 1.01: 1.6; 1.02: 1.5; 1.03: 1.6; 1.04: 1.4856; 1.05: 1.619; 1.06: < 2; 1.07: 1.4645; 1.08: 1.5 |  |  | 1.5036
| 1 | 5 | 1.01: 1.4; 1.02: 1.4; 1.03: 1.6; 1.04: 1.5655; 1.05: 1.621; 1.06: < 2; 1.07: 1.272; 1.08: 1.6 |  |  | 167
| 1 | 6 | 1.01: 1.4; 1.02: 1.5; 1.03: 1.3; 1.04: 1.5975; 1.05: 1.623; 1.06: < 2; 1.07: 1.3895; 1.08: 1.5
| Mean |  | 1.01: 1.4333; 1.02: 1.4667; 1.03: 1.55; 1.04: 1.534; 1.05: 1.6187; 1.07: 1.3724; 1.08: 1.5
| Median |  | 1.01: 1.4; 1.02: 1.5; 1.03: 1.6; 1.04: 1.5275; 1.05: 1.62; 1.07: 1.3737; 1.08: 1.5
| Std Dev. |  | 1.01: 0.10328; 1.02: 0.05164; 1.03: 0.13784; 1.04: 0.043739; 1.05: 0.0081404; 1.07: 0.10548; 1.08: 0.063246
| Rel.Std.Dev. |  | 1.01: 7.206%; 1.02: 3.521%; 1.03: 8.893%; 1.04: 2.851%; 1.05: 0.5029%; 1.07: 7.685%; 1.08: 4.216%
| PDM3 |  | 1.01: -4.673%; 1.02: -2.456%; 1.03: 3.087%; 1.04: 2.025%; 1.05: 7.653%; 1.07: -8.723%; 1.08: -0.2388%
| Z-Score (Absolute) |  | 1.01: 0.57; 1.02: 0.19; 1.03: 0.77; 1.04: 0.58; 1.05: 1.56; 1.06: 5.55; 1.07: 1.27; 1.08: 0.19
| Analytical results for Zn in OREAS 627 (Certified Value 10.04 wt.%) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.15 | 1.16 | 1.17 | 1.19 | 1.20 | 1.23 | 1.25 | 1.26 | 1.27 | 1.28 |  | wt.%
|  |  | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | BF*MS | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS |  | 2
3
| 1 | 1 | 9.36 | 10.1 | 10.4931 | 9.94 | 10.15 | 9.92 | 10.1 | 9.88 | 11 | 10.4249 | 10.15 | 9.5811 | 10.2905 | > 1 | 9.8 | 10.3362 | 9.0224 | 9.93 | 11.45 | 9.913 | 9.95 | 9.89 |  | 1
| 1 | 2 | 9.34 | 10.2 | 10.4701 | 10.05 | 10.3 | 10 | 10 | 10.05 | 10.9 | 10.274 | 9.87 | 9.4328 | 10.2779 | > 1 | 9.8 | 10.3247 | 9.0428 | 9.75 | 11.41 | 9.905 | 10.34 | 9.94 |  | 24
| 1 | 3 | 9.3 | 10.2 | 10.2337 | 10.05 | 10.2 | 10.1 | 10.05 | 10 | 10.9 | 10.1603 | 9.86 | 9.528 | 10.3186 | > 1 | 9.81 | 10.2781 | 9.0658 | 9.9 | 10.35 | 9.921 | 10.16 | 9.88 |  | 16
| 1 | 4 | 9.33 | 10.2 | 10.1663 | 10.05 | 10.2 | 10.25 | 10.2 | 9.96 | 11.6 | 10.3427 | 10.12 | 9.2319 | 10.2686 | > 1 | 9.72 | 10.3266 | 9.009 | 9.8 | 11.35 | 9.917 | 10.16 | 9.92 |  | 10.0404
| 1 | 5 | 9.31 | 10.2 | 10.1254 | 9.98 | 10.25 | 10.15 | 9.98 | 9.86 | 11.6 | 9.9799 | 9.81 | 9.6397 | 10.3159 | > 1 | 9.96 | 10.3285 | 9.0083 | 9.91 | 11.38 | 9.919 | 10.03 | 9.79 |  | 168
| 1 | 6 | 9.32 | 10.2 | 10.3188 | 9.89 | 10.35 | 10.05 | 10.05 | 10.05 | 11.8 | 10.2296 | 9.94 | 9.3233 | 10.2824 | > 1 | 9.62 | 10.3305 | 9.0099 | 9.82 | 11.47 | 9.909 | 10.08 | 9.86
| Mean |  | 9.3267 | 10.1833 | 10.3012 | 9.9933 | 10.2417 | 10.0783 | 10.0633 | 9.9667 | 11.3 | 10.2352 | 9.9583 | 9.4561 | 10.2923 |  | 9.785 | 10.3208 | 9.0264 | 9.8517 | 11.235 | 9.914 | 10.12 | 9.88
| Median |  | 9.325 | 10.2 | 10.2762 | 10.015 | 10.225 | 10.075 | 10.05 | 9.98 | 11.3 | 10.2518 | 9.905 | 9.4804 | 10.2864 |  | 9.8 | 10.3276 | 9.0161 | 9.86 | 11.395 | 9.915 | 10.12 | 9.885
| Std Dev. |  | 0.021602 | 0.040825 | 0.15448 | 0.068313 | 0.073598 | 0.11583 | 0.079162 | 0.082381 | 0.40988 | 0.15483 | 0.14331 | 0.15681 | 0.020587 |  | 0.11238 | 0.021273 | 0.023409 | 0.071949 | 0.43579 | 0.0061644 | 0.13431 | 0.052536
| Rel.Std.Dev. |  | 0.2316% | 0.4009% | 1.5% | 0.6836% | 0.7186% | 1.149% | 0.7866% | 0.8266% | 3.627% | 1.513% | 1.439% | 1.658% | 0.2% |  | 1.149% | 0.2061% | 0.2593% | 0.7303% | 3.879% | 0.06218% | 1.327% | 0.5317%
| PDM3 |  | -7.108% | 1.424% | 2.598% | -0.4684% | 2.005% | 0.3782% | 0.2288% | -0.734% | 12.55% | 1.941% | -0.817% | -5.819% | 2.509% |  | -2.543% | 2.793% | -10.1% | -1.879% | 11.9% | -1.259% | 0.7931% | -1.597%
| Z-Score (Absolute) |  | 2.71 | 0.44 | 0.88 | 0.26 | 0.66 | 0.06 | 0 | 0.36 | 4.55 | 0.63 | 0.39 | 2.23 | 0.84 | NA | 1.02 | 0.95 | 3.81 | 0.78 | 4.31 | 0.55 | 0.21 | 0.67
| Analytical results for Zr in OREAS 627 (Indicative Value 182 ppm) | [24 empty] | Indicative
Zr | Round | Replicate | 1.01: Lab; 1.02: Lab; 1.03: Lab; 1.04: Lab; 1.05: Lab |  | 1
| No | No | 1.01: 1.10; 1.02: 1.13; 1.03: 1.16; 1.04: 1.20; 1.05: 1.23 |  | ppm
|  |  | 1.01: PF*MS; 1.02: PF*OES/MS; 1.03: PF*OES; 1.04: PF*MS; 1.05: PF*MS |  | 0
0
| 1 | 1 | 1.01: 181; 1.02: 184.7777; 1.03: 163.6; 1.04: 177; 1.05: 180 |  | 1
| 1 | 2 | 1.01: 185; 1.02: 181.1737; 1.03: 159.3; 1.04: 177; 1.05: 176 |  | 15
| 1 | 3 | 1.01: 186; 1.02: 182.9531; 1.03: 163.6; 1.04: 167; 1.05: 180 |  | 16
| 1 | 4 | 1.01: 189; 1.02: 180.2532; 1.03: 154.1; 1.04: 182; 1.05: 180 |  | 181.8965
| 1 | 5 | 1.01: 186; 1.02: 187.8597; 1.03: 159.6; 1.04: 185; 1.05: 185 |  | 21
| 1 | 6 | 1.01: 181; 1.02: 188.4993; 1.03: 154.1; 1.04: 174; 1.05: 189
| Mean |  | 1.01: 184.6667; 1.02: 184.2528; 1.03: 159.05; 1.04: 177; 1.05: 181.6667
| Median |  | 1.01: 185.5; 1.02: 183.8654; 1.03: 159.45; 1.04: 177; 1.05: 180
| Std Dev. |  | 1.01: 3.1411; 1.02: 3.4201; 1.03: 4.2609; 1.04: 6.2929; 1.05: 4.5898
| Rel.Std.Dev. |  | 1.01: 1.701%; 1.02: 1.856%; 1.03: 2.679%; 1.04: 3.555%; 1.05: 2.527%
| PDM3 |  | 1.01: 1.523%; 1.02: 1.295%; 1.03: -12.56%; 1.04: -2.692%; 1.05: -0.1264%
| Z-Score (Absolute) |  | 1.01: 0.67; 1.02: 0.58; 1.03: 5.08; 1.04: 1.05; 1.05: 0
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C72: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C90: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C108: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C126: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C144: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C163: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C199: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C217: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C235: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C254: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C272: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C291: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C309: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C327: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C345: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C382: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C400: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C418: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C436: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C454: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C472: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C490: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C509: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C527: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C563: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C581: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C599: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C617: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C635: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C653: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C689: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C707: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C726: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C744: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C762: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C780: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C798: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C816: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C834: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C852: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C871: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C889: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C907: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C925: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C944: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C962: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C981: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C999: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1017: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1035: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1054: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1072: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1090: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1109: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1127: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1145: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for C in OREAS 627 (Indicative Value 0.342 wt.%) | [28 empty] | Indicative
C | Round | Replicate | Lab | Lab | Lab | [24 empty] | 1
| No | No | 1.12 | 1.20 | 3.01 | [24 empty] | wt.%
|  |  | IRC | IRC | IRC | [24 empty] | 3
3
| 1 | 1 | 0.37 | 0.34 | 0.32 | [24 empty] | 1
| 1 | 2 | 0.36 | 0.33 | 0.32 | [24 empty] | 18
| 1 | 3 | 0.36 | 0.34 | [25 empty] | 16
| 1 | 4 | 0.37 | 0.35 | [25 empty] | 0.34167
| 1 | 5 | 0.38 | 0.34 | [25 empty] | 24
| 1 | 6 | 0.34 | 0.35
| Mean |  | 0.36333 | 0.34167 | 0.32
| Median |  | 0.365 | 0.34 | 0.32
| Std Dev. |  | 0.013663 | 0.0075277 | 0
| Rel.Std.Dev. |  | 3.76% | 2.203% | 0%
| PDM3 |  | 6.341% | -1.11e-13% | -6.341%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.67
| Analytical results for S in OREAS 627 (Certified Value 11.05 wt.%) | [28 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 1.01 | 1.02 | 1.03 | 1.04 | 1.05 | 1.06 | 1.07 | 1.08 | 1.09 | 1.10 | 1.11 | 1.12 | 1.13 | 1.14 | 1.15 | 1.16 | 1.17 | 1.18 | 1.19 | 1.20 | 1.21 | 1.23 | 1.24 | 1.25 | 1.26 | 1.27 | 3.01 | wt.%
|  |  | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | IRC | 2
3
| 1 | 1 | 10.83 | 11.6 | 11.7 | 10.75 | 11 | 10.95 | 11 | 10.5 | 12.5 | 11 | 10.98 | 11.13 | 10.8 | 11.03 | 10.6038 | 11.1 | 11.3 | 10.16 | 9.9 | 11.52 | 10.9 | 10.823 | 9.88 | 10.7 | 11.25 | 11.38 | 11.3 | 1
| 1 | 2 | 10.81 | 11.5 | 11.62 | 10.75 | 10.9 | 10.85 | 10.75 | 10.3 | 12.5 | 11.12 | 11.17 | 11.13 | 10.7 | 11.07 | 10.6251 | 10.98 | 11.3 | 10.2 | 9.92 | 11.24 | 11.05 | 10.74 | 9.81 | 10.8 | 11.2 | 11.5 | 11.8 | 26
| 1 | 3 | 10.95 | 11.6 | 11.36 | 10.65 | 11.1 | 10.75 | 10.85 | 10.75 | 12.3 | 11.12 | 10.87 | 11.13 | 10.7 | 11.14 | 10.6233 | 11.13 | 11.3 | 10.19 | 9.95 | 11.49 | 10.92 | 10.823 | 9.85 | 10.1 | 11.25 | 11.54 |  | 16
| 1 | 4 | 10.93 | 11.5 | 11.9 | 10.8 | 11 | 10.75 | 11 | 10.35 | 12.5 | 11.11 | 10.85 | 11.13 | 10.9 | 11.04 | 10.5276 | 10.91 | 11.3 | 10.14 | 9.92 | 11.51 | 10.85 | 10.624 | 9.75 | 10.9 | 11.25 | 11.48 |  | 11.0498
| 1 | 5 | 10.93 | 11.5 | 10.04 | 10.85 | 10.8 | 10.75 | 10.7 | 10.55 | 12.3 | 10.98 | 11.04 | 11.13 | 10.6 | 10.93 | 10.5777 | 10.94 | 11.3 | 10.23 | 9.93 | 11.44 | 10.96 | 10.845 | 9.81 | 10.9 | 11.2 | 11.44 |  | 170
| 1 | 6 | 10.93 | 11.5 | 11.32 | 10.75 | 10.9 | 11.05 | 11.25 | 10.65 | 13.4 | 11.16 | 10.83 | 11.13 | 10.8 | 11.24 | 10.5258 | 11.68 | 11.2 | 10.17 | 9.94 | 11.56 | 10.93 | 10.731 | 9.82 | 10.9 | 11.25 | 11.56
| Mean |  | 10.8967 | 11.5333 | 11.3233 | 10.7583 | 10.95 | 10.85 | 10.925 | 10.5167 | 12.5833 | 11.0817 | 10.9567 | 11.13 | 10.75 | 11.075 | 10.5806 | 11.1233 | 11.2833 | 10.1817 | 9.9267 | 11.46 | 10.935 | 10.7643 | 9.82 | 10.7167 | 11.2333 | 11.4833 | 11.55
| Median |  | 10.93 | 11.5 | 11.49 | 10.75 | 10.95 | 10.8 | 10.925 | 10.525 | 12.5 | 11.115 | 10.925 | 11.13 | 10.75 | 11.055 | 10.5908 | 11.04 | 11.3 | 10.18 | 9.925 | 11.5 | 10.925 | 10.7815 | 9.815 | 10.85 | 11.25 | 11.49 | 11.55
| Std Dev. |  | 0.060222 | 0.05164 | 0.66494 | 0.066458 | 0.10488 | 0.12649 | 0.20187 | 0.17224 | 0.41191 | 0.073326 | 0.13261 | 0 | 0.10488 | 0.10559 | 0.04509 | 0.2864 | 0.040825 | 0.031885 | 0.017512 | 0.11472 | 0.067157 | 0.083411 | 0.043818 | 0.31252 | 0.02582 | 0.066232 | 0.35355
| Rel.Std.Dev. |  | 0.5527% | 0.4477% | 5.872% | 0.6177% | 0.9578% | 1.166% | 1.848% | 1.638% | 3.273% | 0.6617% | 1.21% | 0% | 0.9756% | 0.9534% | 0.4262% | 2.575% | 0.3618% | 0.3132% | 0.1764% | 1.001% | 0.6141% | 0.7749% | 0.4462% | 2.916% | 0.2299% | 0.5768% | 3.061%
| PDM3 |  | -1.386% | 4.376% | 2.476% | -2.638% | -0.9031% | -1.808% | -1.129% | -4.825% | 13.88% | 0.2885% | -0.8428% | 0.7259% | -2.713% | 0.2281% | -4.247% | 0.6655% | 2.114% | -7.856% | -10.16% | 3.712% | -1.039% | -2.583% | -11.13% | -3.015% | 1.661% | 3.924% | 4.527%
| Z-Score (Absolute) |  | 0.15 | 1.69 | 1.08 | 0.55 | 0 | 0.29 | 0.07 | 1.25 | 4.72 | 0.38 | 0.02 | 0.52 | 0.58 | 0.36 | 1.07 | 0.5 | 0.96 | 2.22 | 2.96 | 1.47 | 0.04 | 0.54 | 3.27 | 0.67 | 0.82 | 1.54 | 1.73
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 627 (Indicative Value 10.78 wt.%) |  |  |  | Indicative
Al2O3 | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 10.8135 | 10.73 | 1
| 1 | 2 | 10.8592 | 10.74 | 20
| 1 | 3 | 10.8595 |  | 16
| 1 | 4 | 10.7443 |  | 10.7751
| 1 | 5 | 10.7768 |  | 26
| 1 | 6 | 10.8383
| Mean |  | 10.8153 | 10.735
| Median |  | 10.8259 | 10.735
| Std Dev. |  | 0.046768 | 0.0070711
| Rel.Std.Dev. |  | 0.4324% | 0.06587%
| PDM3 |  | 0.3725% | -0.3725%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for As in OREAS 627 (Indicative Value 9071 ppm) |  |  |  | Indicative
As | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 9064.378 |  | 1
| 1 | 2 | 8961.014 |  | 21
| 1 | 3 | 9110.87 |  | 16
| 1 | 4 | 9037.462 |  | 9071.4902
| 1 | 5 | 9198.623 |  | 27
| 1 | 6 | 9056.594
| Mean |  | 9071.4902
| Median |  | 9060.486
| Std Dev. |  | 79.182
| Rel.Std.Dev. |  | 0.8729%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 627 (Indicative Value 4568 ppm) |  |  |  | Indicative
Ba | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 4510.421 |  | 1
| 1 | 2 | 4546.675 |  | 4
| 1 | 3 | 4671.23 |  | 16
| 1 | 4 | 4613.7 |  | 4568.0553
| 1 | 5 | 4481.298 |  | 28
| 1 | 6 | 4585.008
| Mean |  | 4568.0553
| Median |  | 4565.8415
| Std Dev. |  | 69.7445
| Rel.Std.Dev. |  | 1.527%
| PDM3 |  | 8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 627 (Indicative Value 1.09 wt.%) |  |  |  | Indicative
CaO | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 1.1101 | 1.09 | 1
| 1 | 2 | 1.1131 | 1.08 | 23
| 1 | 3 | 1.1113 |  | 16
| 1 | 4 | 1.0811 |  | 1.0943
| 1 | 5 | 1.1144 |  | 29
| 1 | 6 | 1.0914
| Mean |  | 1.1036 | 1.085
| Median |  | 1.1107 | 1.085
| Std Dev. |  | 0.013882 | 0.0070711
| Rel.Std.Dev. |  | 1.258% | 0.6517%
| PDM3 |  | 0.8483% | -0.8483%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Cr in OREAS 627 (Indicative Value 35.4 ppm) |  |  |  | Indicative
Cr | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 1
1
| 1 | 1 | 23.1476 |  | 1
| 1 | 2 | 33.7977 |  | 15
| 1 | 3 | 75.2796 |  | 16
| 1 | 4 | 9.515 |  | 35.435
| Mean |  | 35.435 |  | 30
| Median |  | 28.4727
| Std Dev. |  | 28.3614
| Rel.Std.Dev. |  | 80.04%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 627 (Indicative Value 0.443 wt.%) |  |  |  | Indicative
Cu | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 3
3
| 1 | 1 | 0.4748 |  | 1
| 1 | 2 | 0.4475 |  | 25
| 1 | 3 | 0.4493 |  | 16
| 1 | 4 | 0.4246 |  | 0.4431
| 1 | 5 | 0.4351 |  | 31
| 1 | 6 | 0.4273
| Mean |  | 0.4431
| Median |  | 0.4413
| Std Dev. |  | 0.01854
| Rel.Std.Dev. |  | 4.184%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe in OREAS 627 (Indicative Value 6.97 wt.%) |  |  |  | Indicative
Fe | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 7.0025 | 6.932 | 1
| 1 | 2 | 7.0087 | 6.98 | 26
| 1 | 3 | 6.947 |  | 16
| 1 | 4 | 6.8959 |  | 6.9653
| 1 | 5 | 6.9841 |  | 32
| 1 | 6 | 7.0096
| Mean |  | 6.9746 | 6.956
| Median |  | 6.9933 | 6.956
| Std Dev. |  | 0.045193 | 0.033941
| Rel.Std.Dev. |  | 0.648% | 0.4879%
| PDM3 |  | 0.1337% | -0.1338%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for K2O in OREAS 627 (Indicative Value 3.1 wt.%) |  |  |  | Indicative
K2O | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 3.1149 | 3.04 | 1
| 1 | 2 | 3.1288 | 3.06 | 27
| 1 | 3 | 3.1611 |  | 16
| 1 | 4 | 3.1369 |  | 3.0973
| 1 | 5 | 3.1761 |  | 33
| 1 | 6 | 3.1503
| Mean |  | 3.1447 | 3.05
| Median |  | 3.1436 | 3.05
| Std Dev. |  | 0.022307 | 0.014142
| Rel.Std.Dev. |  | 0.7093% | 0.4637%
| PDM3 |  | 1.528% | -1.528%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for MgO in OREAS 627 (Indicative Value 0.662 wt.%) |  |  |  | Indicative
MgO | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.6776 | 0.64 | 1
| 1 | 2 | 0.665 | 0.65 | 20
| 1 | 3 | 0.6734 |  | 16
| 1 | 4 | 0.6752 |  | 0.66209
| 1 | 5 | 0.6836 |  | 26
| 1 | 6 | 0.7003
| Mean |  | 0.67918 | 0.645
| Median |  | 0.6764 | 0.645
| Std Dev. |  | 0.011987 | 0.0070711
| Rel.Std.Dev. |  | 1.765% | 1.096%
| PDM3 |  | 2.581% | -2.581%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Mn in OREAS 627 (Indicative Value 0.725 wt.%) |  |  |  | Indicative
Mn | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.7466 | 0.7079 | 1
| 1 | 2 | 0.7368 | 0.7133 | 21
| 1 | 3 | 0.7511 |  | 16
| 1 | 4 | 0.7244 |  | 0.72498
| 1 | 5 | 0.7363 |  | 27
| 1 | 6 | 0.7412
| Mean |  | 0.7394 | 0.7106
| Median |  | 0.739 | 0.7106
| Std Dev. |  | 0.0093096 | 0.0038184
| Rel.Std.Dev. |  | 1.259% | 0.5373%
| PDM3 |  | 1.989% | -1.984%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Na2O in OREAS 627 (Indicative Value 1.61 wt.%) |  |  |  | Indicative
Na2O | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 1.6017 |  | 1
| 1 | 2 | 1.5909 |  | 22
| 1 | 3 | 1.6231 |  | 16
| 1 | 4 | 1.615 |  | 1.605
| 1 | 5 | 1.604 |  | 28
| 1 | 6 | 1.5953
| Mean |  | 1.605
| Median |  | 1.6029
| Std Dev. |  | 0.012103
| Rel.Std.Dev. |  | 0.7541%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 627 (Indicative Value 9.76 ppm) |  |  |  | Indicative
Ni | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 2.0255 |  | 1
| 1 | 2 | 13.9346 |  | 14
| 1 | 3 | 23.2506 |  | 16
| 1 | 4 | 11.3129 |  | 9.7641
| 1 | 5 | 1.0513 |  | 29
| 1 | 6 | 7.0097
| Mean |  | 9.7641
| Median |  | 9.1613
| Std Dev. |  | 8.3074
| Rel.Std.Dev. |  | 85.08%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 627 (Indicative Value 0.037 wt.%) |  |  |  | Indicative
P | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.0358 | 0.0401 | 1
| 1 | 2 | 0.0354 | 0.041 | 24
| 1 | 3 | 0.035 |  | 16
| 1 | 4 | 0.0316 |  | 0.037408
| 1 | 5 | 0.0338 |  | 30
| 1 | 6 | 0.0338
| Mean |  | 0.034233 | 0.04055
| Median |  | 0.0344 | 0.04055
| Std Dev. |  | 0.0015306 | 0.0006364
| Rel.Std.Dev. |  | 4.471% | 1.569%
| PDM3 |  | -8.487% | 8.399%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Pb in OREAS 627 (Indicative Value 2.37 wt.%) |  |  |  | Indicative
Pb | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 2.3678 |  | 1
| 1 | 2 | 2.3551 |  | 25
| 1 | 3 | 2.3739 |  | 16
| 1 | 4 | 2.3466 |  | 2.3665
| 1 | 5 | 2.3842 |  | 31
| 1 | 6 | 2.3712
| Mean |  | 2.3665
| Median |  | 2.3695
| Std Dev. |  | 0.013548
| Rel.Std.Dev. |  | 0.5725%
| PDM3 |  | -1.332e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for S in OREAS 627 (Indicative Value 7.33 wt.%) |  |  |  | Indicative
S | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 3.8202 | 11.0124 | 1
| 1 | 2 | 2.9503 | 10.9723 | 26
| 1 | 3 | 3.8489 |  | 16
| 1 | 4 | 4.6834 |  | 7.3282
| 1 | 5 | 4.1426 |  | 32
| 1 | 6 | 2.5383
| Mean |  | 3.6639 | 10.9924
| Median |  | 3.8346 | 10.9924
| Std Dev. |  | 0.78783 | 0.028355
| Rel.Std.Dev. |  | 21.5% | 0.258%
| PDM3 |  | -50% | 50%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for SiO2 in OREAS 627 (Indicative Value 48.06 wt.%) |  |  |  | Indicative
SiO2 | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 2
2
| 1 | 1 | 48.9348 | 47.08 | 1
| 1 | 2 | 49.0158 | 47.12 | 27
| 1 | 3 | 49.4563 |  | 16
| 1 | 4 | 48.6157 |  | 48.0621
| 1 | 5 | 49.0091 |  | 33
| 1 | 6 | 49.1137
| Mean |  | 49.0242 | 47.1
| Median |  | 49.0125 | 47.1
| Std Dev. |  | 0.27199 | 0.028284
| Rel.Std.Dev. |  | 0.5548% | 0.06005%
| PDM3 |  | 2.002% | -2.002%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Sn in OREAS 627 (Indicative Value 3.77 ppm) |  |  |  | Indicative
Sn | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 3.7702 |  | 1
| Mean |  | 3.7702 |  | 12
| Median |  | 3.7702 |  | 16
| Std Dev. |  |  |  | 3.7702
| Rel.Std.Dev. |  |  |  | 26
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 627 (Indicative Value 71 ppm) |  |  |  | Indicative
Sr | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 147.4688 |  | 1
| 1 | 2 | 133.8941 |  | 21
| 1 | 3 | 62.0112 |  | 16
| 1 | 4 | 13.5833 |  | 71.237
| 1 | 5 | 5.6759 |  | 27
| 1 | 6 | 64.7889
| Mean |  | 71.237
| Median |  | 63.4001
| Std Dev. |  | 59.1373
| Rel.Std.Dev. |  | 83.01%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 627 (Indicative Value 0.246 wt.%) |  |  |  | Indicative
TiO2 | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | wt.%
|  |  | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 0.246 | 0.24 | 1
| 1 | 2 | 0.2508 | 0.241 | 22
| 1 | 3 | 0.2478 |  | 16
| 1 | 4 | 0.2485 |  | 0.24599
| 1 | 5 | 0.2558 |  | 28
| 1 | 6 | 0.26
| Mean |  | 0.25148 | 0.2405
| Median |  | 0.24965 | 0.2405
| Std Dev. |  | 0.0053697 | 0.00070711
| Rel.Std.Dev. |  | 2.135% | 0.294%
| PDM3 |  | 2.232% | -2.232%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for V in OREAS 627 (Indicative Value 19.2 ppm) |  |  |  | Indicative
V | Round | Replicate | BV | BV Geo | 1
| No | No | Ankara | Perth | ppm
|  |  | BF*XRF | BF*XRF | 1
1
| 1 | 1 | 14.6302 | 17 | 1
| 1 | 2 | 26.3544 | 22 | 23
| 1 | 3 | 14.7423 |  | 16
| 1 | 4 | 14.058 |  | 19.2443
| 1 | 5 | 24.6298 |  | 29
| Mean |  | 18.8829 | 19.5
| Median |  | 14.7423 | 19.5
| Std Dev. |  | 6.0696 | 3.5355
| Rel.Std.Dev. |  | 32.14% | 18.13%
| PDM3 |  | -1.878% | 1.329%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Zn in OREAS 627 (Indicative Value 10.55 wt.%) |  |  |  | Indicative
Zn | Round | Replicate | BV |  | 1
| No | No | Ankara |  | wt.%
|  |  | BF*XRF |  | 2
2
| 1 | 1 | 10.6622 |  | 1
| 1 | 2 | 10.6168 |  | 24
| 1 | 3 | 10.3673 |  | 16
| 1 | 4 | 10.5022 |  | 10.5507
| 1 | 5 | 10.5498 |  | 30
| 1 | 6 | 10.6061
| Mean |  | 10.5507
| Median |  | 10.5779
| Std Dev. |  | 0.1057
| Rel.Std.Dev. |  | 1.002%
| PDM3 |  | 3.109e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 627 (Indicative Value 171 ppm) |  |  |  | Indicative
Zr | Round | Replicate | BV |  | 1
| No | No | Ankara |  | ppm
|  |  | BF*XRF |  | 0
0
| 1 | 1 | 203.0663 |  | 1
| 1 | 2 | 182.3528 |  | 15
| 1 | 3 | 136.5913 |  | 16
| 1 | 4 | 150.0325 |  | 171.2761
| 1 | 5 | 187.9562 |  | 31
| 1 | 6 | 167.6573
| Mean |  | 171.2761
| Median |  | 175.0051
| Std Dev. |  | 24.8151
| Rel.Std.Dev. |  | 14.49%
| PDM3 |  | -1.776e-13%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C87: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C105: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C123: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C141: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C159: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C177: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C213: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C231: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C249: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C267: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C285: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C298: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C316: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C334: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C351: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C369: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C387: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 627 (Indicative Value 8.29 wt.%) |  |  | Indicative
LOI1000 | Round | Replicate | BV | 1
| No | No | Ankara | wt.%
|  |  | LOI*TGA | 2
2
| 1 | 1 | 8.27 | 1
| 1 | 2 | 8.34 | 29
| 1 | 3 | 8.27 | 16
| 1 | 4 | 8.34 | 8.29
| 1 | 5 | 8.24 | 35
| 1 | 6 | 8.28
| Mean |  | 8.29
| Median |  | 8.275
| Std Dev. |  | 0.040988
| Rel.Std.Dev. |  | 0.4944%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 627 (Indicative Value 57.5 ppm) |  |  | Indicative
Ag | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 58.6 | 1
| 1 | 2 | 56.3 | 1
| Mean |  | 57.45 | 16
| Median |  | 57.45 | 57.45
| Std Dev. |  | 1.6263 | 37
| Rel.Std.Dev. |  | 2.831%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 627 (Indicative Value 319 ppm) |  |  | Indicative
As | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 318 | 1
| 1 | 2 | 320 | 21
| Mean |  | 319 | 16
| Median |  | 319 | 319
| Std Dev. |  | 1.4142 | 38
| Rel.Std.Dev. |  | 0.4433%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 627 (Indicative Value 4305 ppm) |  |  | Indicative
Ba | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 4320 | 1
| 1 | 2 | 4290 | 4
| Mean |  | 4305 | 16
| Median |  | 4305 | 4305
| Std Dev. |  | 21.2132 | 39
| Rel.Std.Dev. |  | 0.4928%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 627 (Indicative Value 1.8 ppm) |  |  | Indicative
Be | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.6 | 1
| 1 | 2 | 2 | 14
| Mean |  | 1.8 | 16
| Median |  | 1.8 | 1.8
| Std Dev. |  | 0.28284 | 40
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 627 (Indicative Value 45.7 ppm) |  |  | Indicative
Bi | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 46.6 | 1
| 1 | 2 | 44.8 | 35
| Mean |  | 45.7 | 16
| Median |  | 45.7 | 45.7
| Std Dev. |  | 1.2728 | 41
| Rel.Std.Dev. |  | 2.785%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 627 (Indicative Value 66 ppm) |  |  | Indicative
Ce | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 66.2 | 1
| 1 | 2 | 66.6 | 36
| Mean |  | 66.4 | 16
| Median |  | 66.4 | 66.4
| Std Dev. |  | 0.28284 | 42
| Rel.Std.Dev. |  | 0.426%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 627 (Indicative Value 7.85 ppm) |  |  | Indicative
Co | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.9 | 1
| 1 | 2 | 7.8 | 37
| Mean |  | 7.85 | 16
| Median |  | 7.85 | 7.85
| Std Dev. |  | 0.070711 | 43
| Rel.Std.Dev. |  | 0.9008%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 627 (Indicative Value 18 ppm) |  |  | Indicative
Cr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 21 | 1
| 1 | 2 | 15 | 15
| Mean |  | 18 | 16
| Median |  | 18 | 18
| Std Dev. |  | 4.2426 | 44
| Rel.Std.Dev. |  | 23.57%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 627 (Indicative Value 4.97 ppm) |  |  | Indicative
Cs | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.99 | 1
| 1 | 2 | 4.95 | 39
| Mean |  | 4.97 | 16
| Median |  | 4.97 | 4.97
| Std Dev. |  | 0.028284 | 45
| Rel.Std.Dev. |  | 0.5691%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 627 (Indicative Value 0.515 wt.%) |  |  | Indicative
Cu | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.507 | 1
| 1 | 2 | 0.522 | 25
| Mean |  | 0.5145 | 16
| Median |  | 0.5145 | 0.5145
| Std Dev. |  | 0.010607 | 46
| Rel.Std.Dev. |  | 2.062%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 627 (Indicative Value 3.57 ppm) |  |  | Indicative
Dy | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.61 | 1
| 1 | 2 | 3.52 | 7
| Mean |  | 3.565 | 16
| Median |  | 3.565 | 3.565
| Std Dev. |  | 0.06364 | 47
| Rel.Std.Dev. |  | 1.785%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 627 (Indicative Value 1.57 ppm) |  |  | Indicative
Er | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.56 | 1
| 1 | 2 | 1.57 | 8
| Mean |  | 1.565 | 16
| Median |  | 1.565 | 1.565
| Std Dev. |  | 0.0070711 | 48
| Rel.Std.Dev. |  | 0.4518%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 627 (Indicative Value 1.32 ppm) |  |  | Indicative
Eu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.32 | 1
| 1 | 2 | 1.31 | 9
| Mean |  | 1.315 | 16
| Median |  | 1.315 | 1.315
| Std Dev. |  | 0.0070711 | 49
| Rel.Std.Dev. |  | 0.5377%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 627 (Indicative Value 17.9 ppm) |  |  | Indicative
Ga | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.9 | 1
| 1 | 2 | 17.8 | 44
| Mean |  | 17.85 | 16
| Median |  | 17.85 | 17.85
| Std Dev. |  | 0.070711 | 50
| Rel.Std.Dev. |  | 0.3961%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 627 (Indicative Value 4.62 ppm) |  |  | Indicative
Gd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.58 | 1
| 1 | 2 | 4.65 | 10
| Mean |  | 4.615 | 16
| Median |  | 4.615 | 4.615
| Std Dev. |  | 0.049497 | 51
| Rel.Std.Dev. |  | 1.073%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ge in OREAS 627 (Indicative Value 2.05 ppm) |  |  | Indicative
Ge | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.1 | 1
| 1 | 2 | 2 | 46
| Mean |  | 2.05 | 16
| Median |  | 2.05 | 2.05
| Std Dev. |  | 0.070711 | 52
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 627 (Indicative Value 4.91 ppm) |  |  | Indicative
Hf | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.98 | 1
| 1 | 2 | 4.84 | 16
| Mean |  | 4.91 | 16
| Median |  | 4.91 | 4.91
| Std Dev. |  | 0.098995 | 37
| Rel.Std.Dev. |  | 2.016%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 627 (Indicative Value 0.64 ppm) |  |  | Indicative
Ho | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.66 | 1
| 1 | 2 | 0.62 | 6
| Mean |  | 0.64 | 16
| Median |  | 0.64 | 0.64
| Std Dev. |  | 0.028284 | 38
| Rel.Std.Dev. |  | 4.419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 627 (Indicative Value 8.13 ppm) |  |  | Indicative
In | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.2 | 1
| 1 | 2 | 8.05 | 33
| Mean |  | 8.125 | 16
| Median |  | 8.125 | 8.125
| Std Dev. |  | 0.10607 | 39
| Rel.Std.Dev. |  | 1.305%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 627 (Indicative Value 34.3 ppm) |  |  | Indicative
La | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 34.6 | 1
| 1 | 2 | 34 | 34
| Mean |  | 34.3 | 16
| Median |  | 34.3 | 34.3
| Std Dev. |  | 0.42426 | 40
| Rel.Std.Dev. |  | 1.237%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 627 (Indicative Value 0.21 ppm) |  |  | Indicative
Lu | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.21 | 1
| 1 | 2 | 0.2 | 35
| Mean |  | 0.205 | 16
| Median |  | 0.205 | 0.205
| Std Dev. |  | 0.0070711 | 41
| Rel.Std.Dev. |  | 3.449%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 627 (Indicative Value 0.724 wt.%) |  |  | Indicative
Mn | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.725 | 1
| 1 | 2 | 0.722 | 21
| Mean |  | 0.7235 | 16
| Median |  | 0.7235 | 0.7235
| Std Dev. |  | 0.0021213 | 42
| Rel.Std.Dev. |  | 0.2932%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 627 (Indicative Value 30.4 ppm) |  |  | Indicative
Mo | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 31 | 1
| 1 | 2 | 29.8 | 37
| Mean |  | 30.4 | 16
| Median |  | 30.4 | 30.4
| Std Dev. |  | 0.84853 | 43
| Rel.Std.Dev. |  | 2.791%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 627 (Indicative Value 11.1 ppm) |  |  | Indicative
Nb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.3 | 1
| 1 | 2 | 10.9 | 38
| Mean |  | 11.1 | 16
| Median |  | 11.1 | 11.1
| Std Dev. |  | 0.28284 | 44
| Rel.Std.Dev. |  | 2.548%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 627 (Indicative Value 30.9 ppm) |  |  | Indicative
Nd | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30.9 | 1
| 1 | 2 | 30.8 | 7
| Mean |  | 30.85 | 16
| Median |  | 30.85 | 30.85
| Std Dev. |  | 0.070711 | 45
| Rel.Std.Dev. |  | 0.2292%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 627 (Indicative Value 15 ppm) |  |  | Indicative
Ni | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 10 | 14
| Mean |  | 15 | 16
| Median |  | 15 | 15
| Std Dev. |  | 7.0711 | 46
| Rel.Std.Dev. |  | 47.14%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 627 (Indicative Value 2.09 wt.%) |  |  | Indicative
Pb | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.09 | 1
| 1 | 2 | 2.09 | 25
| Mean |  | 2.09 | 16
| Median |  | 2.09 | 2.09
| Std Dev. |  | 0 | 47
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 627 (Indicative Value 8.31 ppm) |  |  | Indicative
Pr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.27 | 1
| 1 | 2 | 8.35 | 9
| Mean |  | 8.31 | 16
| Median |  | 8.31 | 8.31
| Std Dev. |  | 0.056569 | 48
| Rel.Std.Dev. |  | 0.6807%
| PDM3 |  | -2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 627 (Indicative Value 125 ppm) |  |  | Indicative
Rb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 124 | 1
| 1 | 2 | 125 | 43
| Mean |  | 124.5 | 16
| Median |  | 124.5 | 124.5
| Std Dev. |  | 0.70711 | 49
| Rel.Std.Dev. |  | 0.568%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 627 (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 15
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 50
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 627 (Indicative Value 152 ppm) |  |  | Indicative
Sb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 154 | 1
| 1 | 2 | 150 | 45
| Mean |  | 152 | 16
| Median |  | 152 | 152
| Std Dev. |  | 2.8284 | 51
| Rel.Std.Dev. |  | 1.861%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 627 (Indicative Value 5.7 ppm) |  |  | Indicative
Sc | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.9 | 1
| 1 | 2 | 5.5 | 46
| Mean |  | 5.7 | 16
| Median |  | 5.7 | 5.7
| Std Dev. |  | 0.28284 | 52
| Rel.Std.Dev. |  | 4.962%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 627 (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 16
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 37
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 627 (Indicative Value 5.86 ppm) |  |  | Indicative
Sm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.8 | 1
| 1 | 2 | 5.91 | 8
| Mean |  | 5.855 | 16
| Median |  | 5.855 | 5.855
| Std Dev. |  | 0.077782 | 38
| Rel.Std.Dev. |  | 1.328%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 627 (Indicative Value 6.4 ppm) |  |  | Indicative
Sn | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.2 | 1
| 1 | 2 | 6.6 | 12
| Mean |  | 6.4 | 16
| Median |  | 6.4 | 6.4
| Std Dev. |  | 0.28284 | 39
| Rel.Std.Dev. |  | 4.419%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 627 (Indicative Value 102 ppm) |  |  | Indicative
Sr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 103 | 1
| 1 | 2 | 101 | 21
| Mean |  | 102 | 16
| Median |  | 102 | 102
| Std Dev. |  | 1.4142 | 40
| Rel.Std.Dev. |  | 1.386%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 627 (Indicative Value 0.81 ppm) |  |  | Indicative
Ta | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.78 | 13
| Mean |  | 0.81 | 16
| Median |  | 0.81 | 0.81
| Std Dev. |  | 0.042426 | 41
| Rel.Std.Dev. |  | 5.238%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 627 (Indicative Value 0.66 ppm) |  |  | Indicative
Tb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.65 | 1
| 1 | 2 | 0.67 | 36
| Mean |  | 0.66 | 16
| Median |  | 0.66 | 0.66
| Std Dev. |  | 0.014142 | 42
| Rel.Std.Dev. |  | 2.143%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 627 (Indicative Value 0.25 ppm) |  |  | Indicative
Te | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | < 0.2 | 14
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.25
| Std Dev. |  |  | 43
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 627 (Indicative Value 11.8 ppm) |  |  | Indicative
Th | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 11.6 | 1
| 1 | 2 | 11.9 | 38
| Mean |  | 11.75 | 16
| Median |  | 11.75 | 11.75
| Std Dev. |  | 0.21213 | 44
| Rel.Std.Dev. |  | 1.805%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 627 (Indicative Value 0.156 wt.%) |  |  | Indicative
Ti | Round | Replicate | BV Geo | 1
| No | No | Perth | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.157 | 1
| 1 | 2 | 0.154 | 39
| Mean |  | 0.1555 | 16
| Median |  | 0.1555 | 0.1555
| Std Dev. |  | 0.0021213 | 45
| Rel.Std.Dev. |  | 1.364%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 627 (Indicative Value 0.22 ppm) |  |  | Indicative
Tm | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.23 | 1
| 1 | 2 | 0.21 | 9
| Mean |  | 0.22 | 16
| Median |  | 0.22 | 0.22
| Std Dev. |  | 0.014142 | 46
| Rel.Std.Dev. |  | 6.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 627 (Indicative Value 4.72 ppm) |  |  | Indicative
U | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.8 | 1
| 1 | 2 | 4.64 | 41
| Mean |  | 4.72 | 16
| Median |  | 4.72 | 4.72
| Std Dev. |  | 0.11314 | 47
| Rel.Std.Dev. |  | 2.397%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 627 (Indicative Value 23 ppm) |  |  | Indicative
V | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 23.3 | 1
| 1 | 2 | 22.7 | 23
| Mean |  | 23 | 16
| Median |  | 23 | 23
| Std Dev. |  | 0.42426 | 48
| Rel.Std.Dev. |  | 1.845%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 627 (Indicative Value 8.25 ppm) |  |  | Indicative
W | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.5 | 1
| 1 | 2 | 8 | 43
| Mean |  | 8.25 | 16
| Median |  | 8.25 | 8.25
| Std Dev. |  | 0.35355 | 49
| Rel.Std.Dev. |  | 4.285%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 627 (Indicative Value 18 ppm) |  |  | Indicative
Y | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18 | 1
| 1 | 2 | 18 | 44
| Mean |  | 18 | 16
| Median |  | 18 | 18
| Std Dev. |  | 0 | 50
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 627 (Indicative Value 1.53 ppm) |  |  | Indicative
Yb | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.52 | 45
| Mean |  | 1.525 | 16
| Median |  | 1.525 | 1.525
| Std Dev. |  | 0.0070711 | 51
| Rel.Std.Dev. |  | 0.4637%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 627 (Indicative Value 173 ppm) |  |  | Indicative
Zr | Round | Replicate | BV Geo | 1
| No | No | Perth | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 175 | 1
| 1 | 2 | 171 | 15
| Mean |  | 173 | 16
| Median |  | 173 | 173
| Std Dev. |  | 2.8284 | 52
| Rel.Std.Dev. |  | 1.635%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.