## Abbreviations
Table 1. Abbreviations used for OREAS 462
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 462
Laboratory Name and Location
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
Nagrom, Perth, WA, Australia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
SGS Vostok Limited, Chita, Russian Federation
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 462
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate Fusion XRF
Ce, Cerium (ppm) | 5032.3782 | 123.1831 | 4928.4074 | 5136.3491 | 4874.2642 | 5190.4923
Fe, Iron (wt.%) | 34.0591 | 0.52618 | 33.3452 | 34.7731 | 33.8861 | 34.2322
La, Lanthanum (ppm) | 3851.7898 | 78.2798 | 3787.9724 | 3915.6071 | 3732.5022 | 3971.0774
Nd, Neodymium (ppm) | 2612.8578 | 64.3055 | 2554.3043 | 2671.4113 | 2524.373 | 2701.3426
Pr, Praseodymium (ppm) | 699.5916 | 98.6606 | 591.094 | 808.0891 | IND | IND
Sm, Samarium (ppm) | < 400 | IND | IND | IND | IND | IND
Thermogravimetry
LOI, Loss On Ignition @1000°C (wt.%) | 0.77067 | 0.12029 | 0.60958 | 0.93175 | 0.73666 | 0.80467
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 5.7942 | 0.19525 | 5.6549 | 5.9334 | 5.6803 | 5.908
Ba, Barium (ppm) | 1024.7736 | 52.3375 | 996.1474 | 1053.3999 | 997.9395 | 1051.6078
Be, Beryllium (ppm) | 3.756 | 0.32718 | 3.6083 | 3.9037 | IND | IND
Bi, Bismuth (ppm) | 2.5837 | 0.42071 | 2.126 | 3.0414 | IND | IND
Ca, Calcium (wt.%) | 1.2785 | 0.064108 | 1.237 | 1.3199 | 1.2504 | 1.3065
Ce, Cerium (ppm) | 4951.1664 | 208.022 | 4847.9359 | 5054.397 | 4798.6263 | 5103.7065
Co, Cobalt (ppm) | 12.4505 | 2.0187 | 10.3523 | 14.5487 | 11.7222 | 13.1788
Cr, Chromium (ppm) | 608.7018 | 31.0722 | 590.882 | 626.5215 | 591.6293 | 625.7742
Cs, Cesium (ppm) | 0.43692 | 0.066297 | 0.39173 | 0.4821 | 0.39581 | 0.47803
Dy, Dysprosium (ppm) | 51.1987 | 2.2479 | 49.9278 | 52.4696 | 49.8622 | 52.5352
Er, Erbium (ppm) | 12.2186 | 0.55948 | 11.934 | 12.5032 | 11.7616 | 12.6756
Eu, Europium (ppm) | 78.7127 | 2.8232 | 77.2092 | 80.2162 | 76.5511 | 80.8744
Fe, Iron (wt.%) | 33.836 | 1.4106 | 32.8243 | 34.8476 | 33.0152 | 34.6567
Ga, Gallium (ppm) | 52.2567 | 8.5606 | 42.7906 | 61.7228 | 49.863 | 54.6504
Gd, Gadolinium (ppm) | 165.2292 | 7.3324 | 161.427 | 169.0315 | 160.6232 | 169.8352
Hf, Hafnium (ppm) | 13.9638 | 0.7841 | 13.6719 | 14.2558 | 13.3036 | 14.6241
Ho, Holmium (ppm) | 6.4951 | 0.46035 | 6.2459 | 6.7443 | 6.1871 | 6.8032
K, Potassium (wt.%) | 0.12025 | 0.017123 | 0.10613 | 0.13436 | IND | IND
La, Lanthanum (ppm) | 3794.3497 | 73.3896 | 3765.0828 | 3823.6167 | 3697.5141 | 3891.1853
Li, Lithium (ppm) | 11.3006 | 0.84294 | 10.3136 | 12.2877 | IND | IND
Lu, Lutetium (ppm) | 0.62929 | 0.043245 | 0.6074 | 0.65118 | 0.57585 | 0.68274
Mg, Magnesium (wt.%) | 1.0831 | 0.03186 | 1.0594 | 1.1067 | 1.0631 | 1.1031
Mn, Manganese (ppm) | 932.7236 | 67.5368 | 880.85 | 984.5971 | 912.9749 | 952.4723
Mo, Molybdenum (ppm) | 51.4559 | 3.2528 | 48.8681 | 54.0436 | 48.9569 | 53.9549
Nb, Niobium (ppm) | 1447.3519 | 38.9739 | 1421.5273 | 1473.1764 | 1404.0438 | 1490.6599
Nd, Neodymium (ppm) | 2559.577 | 127.8725 | 2487.259 | 2631.895 | 2485.5604 | 2633.5935
Ni, Nickel (ppm) | 66.4512 | 8.9966 | 58.8524 | 74.05 | 51.6695 | 81.2329
P, Phosphorus (wt.%) | 0.49237 | 0.010433 | 0.48554 | 0.49919 | 0.47859 | 0.50614
Pb, Lead (ppm) | 114.3628 | 5.6555 | 109.1004 | 119.6253 | 109.5695 | 119.1562
Pr, Praseodymium (ppm) | 735.2301 | 33.7476 | 718.4125 | 752.0477 | 716.1686 | 754.2915
Rb, Rubidium (ppm) | 6.577 | 0.37655 | 6.4037 | 6.7502 | 6.258 | 6.8959
S, Sulphur (ppm) | 587 | 88.5574 | 523.0747 | 650.9253 | IND | IND
Sb, Antimony (ppm) | 2.5686 | 0.38164 | 2.2852 | 2.8521 | 2.2661 | 2.8712
Si, Silicon (wt.%) | 13.1762 | 0.24877 | 12.9687 | 13.3838 | 12.9502 | 13.4022
Sm, Samarium (ppm) | 368.837 | 8.8893 | 365.1297 | 372.5443 | 358.4841 | 379.1898
Sn, Tin (ppm) | 29.2167 | 3.304 | 27.2165 | 31.2169 | 27.3773 | 31.0561
Sr, Strontium (ppm) | 764.6318 | 21.5656 | 754.4008 | 774.8629 | 742.0482 | 787.2154
Ta, Tantalum (ppm) | 25.6598 | 1.6503 | 24.6371 | 26.6824 | 24.0725 | 27.247
Tb, Terbium (ppm) | 14.3935 | 0.53793 | 14.1259 | 14.6611 | 13.9511 | 14.8359
Th, Thorium (ppm) | 249.0714 | 8.9058 | 244.4954 | 253.6474 | 242.5223 | 255.6205
Ti, Titanium (wt.%) | 1.9075 | 0.073938 | 1.8593 | 1.9558 | 1.8646 | 1.9504
Tm, Thulium (ppm) | 1.2039 | 0.10316 | 1.1549 | 1.2529 | 1.1034 | 1.3043
U, Uranium (ppm) | 6.1476 | 0.26636 | 6.0275 | 6.2677 | 5.7968 | 6.4984
V, Vanadium (ppm) | 380.1715 | 18.0481 | 371.7276 | 388.6153 | 364.0047 | 396.3383
W, Tungsten (ppm) | 3.5952 | 0.66478 | 3.055 | 4.1355 | IND | IND
Y, Yttrium (ppm) | 132.6531 | 6.4155 | 129.0046 | 136.3016 | 129.1054 | 136.2008
Yb, Ytterbium (ppm) | 5.5122 | 0.36102 | 5.341 | 5.6833 | 5.2442 | 5.7801
Zn, Zinc (ppm) | 268.9 | 25.9575 | 233.4812 | 304.3188 | IND | IND
Zr, Zirconium (ppm) | 593.3136 | 26.8707 | 577.0266 | 609.6006 | 569.2937 | 617.3335
4-Acid Digestion
Ag, Silver (ppm) | < 2 | IND | IND | IND | IND | IND
Al, Aluminium (wt.%) | 5.7069 | 0.13173 | 5.6261 | 5.7877 | 5.6006 | 5.8132
As, Arsenic (ppm) | 30.7827 | 1.2784 | 29.9043 | 31.661 | 29.4099 | 32.1555
Ba, Barium (ppm) | 1036.7978 | 37.6574 | 1014.4999 | 1059.0957 | 1016.0366 | 1057.559
Be, Beryllium (ppm) | 3.5673 | 0.32448 | 3.3647 | 3.77 | 3.4137 | 3.721
Bi, Bismuth (ppm) | 2.6935 | 0.17256 | 2.5927 | 2.7943 | 2.5671 | 2.8199
Ca, Calcium (wt.%) | 1.2475 | 0.08679 | 1.1911 | 1.304 | 1.2244 | 1.2707
Cd, Cadmium (ppm) | < 0.2 | IND | IND | IND | IND | IND
Ce, Cerium (ppm) | 4984.2902 | 180.3861 | 4827.0384 | 5141.542 | 4897.4857 | 5071.0947
Co, Cobalt (ppm) | 13.0237 | 0.87369 | 12.4815 | 13.5659 | 12.6886 | 13.3588
Cr, Chromium (ppm) | 501.5086 | 81.5901 | 444.8845 | 558.1327 | 486.2858 | 516.7314
Cs, Cesium (ppm) | 0.45029 | 0.063532 | 0.41437 | 0.48621 | 0.42519 | 0.47539
Cu, Copper (ppm) | 60.8857 | 1.942 | 59.5775 | 62.1939 | 59.3652 | 62.4061
Dy, Dysprosium (ppm) | 50.5936 | 1.2121 | 50.0287 | 51.1585 | 49.556 | 51.6312
Er, Erbium (ppm) | 10.8497 | 0.4193 | 10.5928 | 11.1066 | 10.5431 | 11.1564
Eu, Europium (ppm) | 79.8818 | 4.4042 | 77.2148 | 82.5487 | 77.9651 | 81.7985
Fe, Iron (wt.%) | 33.0543 | 0.99235 | 32.4331 | 33.6754 | 32.5313 | 33.5772
Ga, Gallium (ppm) | 58.3472 | 7.0937 | 48.8694 | 67.825 | 55.9011 | 60.7933
Gd, Gadolinium (ppm) | 163.8761 | 5.5416 | 160.5969 | 167.1552 | 160.1515 | 167.6006
Hf, Hafnium (ppm) | 7.3979 | 0.83226 | 6.7892 | 8.0066 | 6.9995 | 7.7963
Ho, Holmium (ppm) | 6.1793 | 0.32344 | 5.9802 | 6.3784 | 5.9863 | 6.3723
In, Indium (ppm) | 0.74891 | 0.038885 | 0.72688 | 0.77093 | 0.72514 | 0.77267
K, Potassium (wt.%) | 0.11843 | 0.0053865 | 0.11517 | 0.12169 | IND | IND
La, Lanthanum (ppm) | 3669.026 | 193.4809 | 3516.4667 | 3821.5853 | 3577.6518 | 3760.4002
Li, Lithium (ppm) | 10.8383 | 1.4008 | 10.0776 | 11.5991 | 10.4053 | 11.2714
Lu, Lutetium (ppm) | 0.52948 | 0.096701 | 0.47114 | 0.58781 | 0.50008 | 0.55887
Mg, Magnesium (wt.%) | 1.007 | 0.11892 | 0.93371 | 1.0803 | 0.9873 | 1.0267
Mn, Manganese (ppm) | 839.9573 | 63.328 | 789.8353 | 890.0792 | 819.8507 | 860.0638
Mo, Molybdenum (ppm) | 53.4403 | 3.5067 | 51.1924 | 55.6881 | 52.2595 | 54.6211
Na, Sodium (wt.%) | 0.17785 | 0.012132 | 0.17063 | 0.18506 | IND | IND
Nd, Neodymium (ppm) | 2595.2838 | 125.645 | 2522.5226 | 2668.0449 | 2534.5411 | 2656.0265
Ni, Nickel (ppm) | 73.8248 | 3.1147 | 71.9218 | 75.7277 | 71.0036 | 76.6459
P, Phosphorus (wt.%) | 0.45372 | 0.049008 | 0.42287 | 0.48456 | 0.44279 | 0.46465
Pb, Lead (ppm) | 118.5188 | 5.3267 | 114.6637 | 122.374 | 116.1455 | 120.8922
Pr, Praseodymium (ppm) | 735.9837 | 29.8461 | 720.3896 | 751.5778 | 719.8265 | 752.141
Rb, Rubidium (ppm) | 6.4463 | 0.60099 | 6.1153 | 6.7774 | 6.2015 | 6.6911
Re, Rhenium (ppm) | < 0.01 | IND | IND | IND | IND | IND
S, Sulphur (ppm) | 547.9444 | 70.186 | 503.5351 | 592.3538 | IND | IND
Sb, Antimony (ppm) | 2.172 | 0.2516 | 2.0062 | 2.3379 | 2.0568 | 2.2873
Sc, Scandium (ppm) | 51.4401 | 3.7349 | 49.1884 | 53.6918 | 49.6775 | 53.2027
Sm, Samarium (ppm) | 361.9022 | 17.3861 | 352.1735 | 371.6309 | 352.9464 | 370.858
Sn, Tin (ppm) | 23.8764 | 1.3741 | 22.8764 | 24.8765 | 22.9797 | 24.7731
Sr, Strontium (ppm) | 737.5233 | 58.5207 | 700.6018 | 774.4448 | 727.2487 | 747.798
Ta, Tantalum (ppm) | 21.8115 | 2.3785 | 19.6885 | 23.9344 | 20.8779 | 22.7451
Tb, Terbium (ppm) | 14.7025 | 0.62042 | 14.3337 | 15.0713 | 14.2834 | 15.1215
Te, Tellurium (ppm) | 0.33312 | 0.061006 | 0.27671 | 0.38953 | 0.2929 | 0.37333
Th, Thorium (ppm) | 245.568 | 7.9044 | 241.1112 | 250.0247 | 237.4841 | 253.6518
Ti, Titanium (wt.%) | 0.98596 | 0.10432 | 0.90759 | 1.0643 | 0.93448 | 1.0374
Tl, Thallium (ppm) | 0.082847 | 0.010016 | 0.076054 | 0.089639 | IND | IND
Tm, Thulium (ppm) | 1.011 | 0.089954 | 0.95899 | 1.0631 | 0.96231 | 1.0598
U, Uranium (ppm) | 5.7922 | 0.36416 | 5.5959 | 5.9886 | 5.5247 | 6.0598
V, Vanadium (ppm) | 352.6418 | 26.0486 | 333.5665 | 371.717 | 347.3162 | 357.9673
W, Tungsten (ppm) | 2.3069 | 0.33121 | 2.0693 | 2.5444 | 2.1501 | 2.4636
Y, Yttrium (ppm) | 125.2743 | 4.7867 | 122.2897 | 128.2589 | 122.1398 | 128.4088
Yb, Ytterbium (ppm) | 4.3644 | 0.35244 | 4.1431 | 4.5856 | 4.1951 | 4.5336
Zn, Zinc (ppm) | 243.663 | 24.7633 | 226.6817 | 260.6443 | 239.2735 | 248.0525
Zr, Zirconium (ppm) | 270.2672 | 30.9223 | 249.3799 | 291.1545 | 260.0266 | 280.5078

## Indicative Values
Table 3. Indicative Values for OREAS 462
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Al | wt.% | 5.8669 | Lu | ppm | < 20 | Tb | ppm | 17.0039
Ba | ppm | 1119.5701 | Mg | wt.% | 1.0687 | Th | ppm | 261.7139
Ca | wt.% | 1.3135 | Mn | ppm | 955.1838 | Ti | wt.% | 1.9139
Cr | ppm | 423.0813 | Na | wt.% | 0.19071 | Tm | ppm | < 10
Dy | ppm | 52.2785 | Nb | ppm | 1452.3739 | U | ppm | < 100
Er | ppm | 17.4902 | P | wt.% | 0.50086 | V | ppm | 417.3202
Eu | ppm | 71.9694 | Si | wt.% | 13.334 | W | ppm | < 100
Gd | ppm | 162.6757 | Sn | ppm | < 100 | Y | ppm | 88.4558
Hf | ppm | < 100 | S | wt.% | 0.055662 | Yb | ppm | 8.7819
Ho | ppm | 8.7298 | Sr | ppm | 620.1026 | Zr | ppm | 530.5498
K | wt.% | 0.12397 | Ta | ppm | < 100
Thermogravimetry
H2O- | wt.% | 0.66667
Borate / Peroxide Fusion ICP
Ag | ppm | 9.5 | Ge | ppm | 2.6667 | Se | ppm | < 20
As | ppm | 99.1665 | In | ppm | 0.78261 | Te | ppm | < 1
B | ppm | 150.25 | Na | wt.% | 0.16744 | Tl | ppm | < 0.5
Cd | ppm | < 1 | Re | ppm | < 0.1
Cu | ppm | 72.3442 | Sc | ppm | 74.25
4-Acid Digestion
Ge | ppm | 6 | Nb | ppm | 1231.0306 | Se | ppm | 8.2354

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 462
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Ce, ppm | 5032.3782 | 123.1831 | 4786.0121 | 5278.7444 | 4662.829 | 5401.9275 | 2.448% | 4.896% | 7.343% | 4780.7593 | 5283.9972
Fe, wt.% | 34.0591 | 0.52618 | 33.0068 | 35.1115 | 32.4806 | 35.6377 | 1.545% | 3.09% | 4.635% | 32.3562 | 35.7621
La, ppm | 3851.7898 | 78.2798 | 3695.2302 | 4008.3494 | 3616.9504 | 4086.6291 | 2.032% | 4.065% | 6.097% | 3659.2003 | 4044.3793
Nd, ppm | 2612.8578 | 64.3055 | 2484.2467 | 2741.4689 | 2419.9412 | 2805.7744 | 2.461% | 4.922% | 7.383% | 2482.2149 | 2743.5007
Pr, ppm | 699.5916 | 98.6606 | 502.2704 | 896.9127 | 403.6098 | 995.5733 | 14.1% | 28.21% | 42.31% | 664.612 | 734.5711
Sm, ppm | < 400 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Thermogravimetry
LOI, wt.% | 0.77067 | 0.12029 | 0.53008 | 1.0113 | 0.40979 | 1.1315 | 15.61% | 31.22% | 46.83% | 0.73213 | 0.8092
Borate / Peroxide Fusion ICP
Al, wt.% | 5.7942 | 0.19525 | 5.4037 | 6.1847 | 5.2084 | 6.3799 | 3.37% | 6.739% | 10.11% | 5.5045 | 6.0839
Ba, ppm | 1024.7736 | 52.3375 | 920.0987 | 1129.4486 | 867.7612 | 1181.7861 | 5.107% | 10.21% | 15.32% | 973.535 | 1076.0123
Be, ppm | 3.756 | 0.32718 | 3.1016 | 4.4104 | 2.7745 | 4.7375 | 8.711% | 17.42% | 26.13% | 3.5682 | 3.9438
Bi, ppm | 2.5837 | 0.42071 | 1.7423 | 3.4251 | 1.3216 | 3.8458 | 16.28% | 32.57% | 48.85% | 2.4545 | 2.7129
Ca, wt.% | 1.2785 | 0.064108 | 1.1502 | 1.4067 | 1.0861 | 1.4708 | 5.015% | 10.03% | 15.04% | 1.2145 | 1.3424
Ce, ppm | 4951.1664 | 208.022 | 4535.1224 | 5367.2104 | 4327.1004 | 5575.2324 | 4.201% | 8.403% | 12.6% | 4703.6081 | 5198.7247
Co, ppm | 12.4505 | 2.0187 | 8.4131 | 16.488 | 6.3943 | 18.5067 | 16.21% | 32.43% | 48.64% | 11.828 | 13.0731
Cr, ppm | 608.7018 | 31.0722 | 546.5574 | 670.8461 | 515.4852 | 701.9183 | 5.105% | 10.21% | 15.31% | 578.2667 | 639.1368
Cs, ppm | 0.43692 | 0.066297 | 0.30432 | 0.56951 | 0.23803 | 0.63581 | 15.17% | 30.35% | 45.52% | 0.41507 | 0.45876
Dy, ppm | 51.1987 | 2.2479 | 46.7028 | 55.6946 | 44.4549 | 57.9425 | 4.391% | 8.781% | 13.17% | 48.6388 | 53.7586
Er, ppm | 12.2186 | 0.55948 | 11.0996 | 13.3376 | 10.5402 | 13.8971 | 4.579% | 9.158% | 13.74% | 11.6077 | 12.8295
Eu, ppm | 78.7127 | 2.8232 | 73.0664 | 84.359 | 70.2432 | 87.1822 | 3.587% | 7.173% | 10.76% | 74.7771 | 82.6484
Fe, wt.% | 33.836 | 1.4106 | 31.0147 | 36.6572 | 29.6041 | 38.0678 | 4.169% | 8.338% | 12.51% | 32.1442 | 35.5278
Ga, ppm | 52.2567 | 8.5606 | 35.1354 | 69.378 | 26.5748 | 77.9386 | 16.38% | 32.76% | 49.15% | 49.6439 | 54.8696
Gd, ppm | 165.2292 | 7.3324 | 150.5645 | 179.894 | 143.2321 | 187.2263 | 4.438% | 8.875% | 13.31% | 156.9678 | 173.4907
Hf, ppm | 13.9638 | 0.7841 | 12.3956 | 15.532 | 11.6115 | 16.3161 | 5.615% | 11.23% | 16.85% | 13.2656 | 14.662
Ho, ppm | 6.4951 | 0.46035 | 5.5744 | 7.4158 | 5.1141 | 7.8762 | 7.088% | 14.18% | 21.26% | 6.1704 | 6.8199
K, wt.% | 0.12025 | 0.017123 | 0.086001 | 0.15449 | 0.068878 | 0.17162 | 14.24% | 28.48% | 42.72% | 0.11424 | 0.12626
La, ppm | 3794.3497 | 73.3896 | 3647.5705 | 3941.129 | 3574.1808 | 4014.5187 | 1.934% | 3.868% | 5.803% | 3604.6323 | 3984.0672
Li, ppm | 11.3006 | 0.84294 | 9.6148 | 12.9865 | 8.7718 | 13.8295 | 7.459% | 14.92% | 22.38% | 10.7356 | 11.8657
Lu, ppm | 0.62929 | 0.043245 | 0.5428 | 0.71578 | 0.49956 | 0.75903 | 6.872% | 13.74% | 20.62% | 0.59783 | 0.66076
Mg, wt.% | 1.0831 | 0.03186 | 1.0193 | 1.1468 | 0.98748 | 1.1786 | 2.942% | 5.883% | 8.825% | 1.0289 | 1.1372
Mn, ppm | 932.7236 | 67.5368 | 797.65 | 1067.7971 | 730.1133 | 1135.3339 | 7.241% | 14.48% | 21.72% | 886.0874 | 979.3598
Mo, ppm | 51.4559 | 3.2528 | 44.9503 | 57.9614 | 41.6976 | 61.2142 | 6.321% | 12.64% | 18.96% | 48.8831 | 54.0286
Nb, ppm | 1447.3519 | 38.9739 | 1369.404 | 1525.2997 | 1330.4301 | 1564.2736 | 2.693% | 5.386% | 8.078% | 1374.9843 | 1519.7194
Nd, ppm | 2559.577 | 127.8725 | 2303.832 | 2815.322 | 2175.9595 | 2943.1945 | 4.996% | 9.992% | 14.99% | 2431.5981 | 2687.5558
Ni, ppm | 66.4512 | 8.9966 | 48.4581 | 84.4443 | 39.4615 | 93.4408 | 13.54% | 27.08% | 40.62% | 63.1286 | 69.7738
P, wt.% | 0.49237 | 0.010433 | 0.4715 | 0.51323 | 0.46107 | 0.52366 | 2.119% | 4.238% | 6.357% | 0.46775 | 0.51698
Pb, ppm | 114.3628 | 5.6555 | 103.0519 | 125.6738 | 97.3964 | 131.3293 | 4.945% | 9.89% | 14.84% | 108.6447 | 120.081
Pr, ppm | 735.2301 | 33.7476 | 667.7348 | 802.7253 | 633.9872 | 836.4729 | 4.59% | 9.18% | 13.77% | 698.4686 | 771.9916
Rb, ppm | 6.577 | 0.37655 | 5.8238 | 7.3301 | 5.4473 | 7.7066 | 5.725% | 11.45% | 17.18% | 6.2481 | 6.9058
S, ppm | 587 | 88.5574 | 409.8852 | 764.1148 | 321.3278 | 852.6722 | 15.09% | 30.17% | 45.26% | 557.65 | 616.35
Sb, ppm | 2.5686 | 0.38164 | 1.8054 | 3.3319 | 1.4237 | 3.7136 | 14.86% | 29.72% | 44.57% | 2.4402 | 2.6971
Si, wt.% | 13.1762 | 0.24877 | 12.6787 | 13.6737 | 12.4299 | 13.9225 | 1.888% | 3.776% | 5.664% | 12.5174 | 13.835
Sm, ppm | 368.837 | 8.8893 | 351.0584 | 386.6155 | 342.1691 | 395.5048 | 2.41% | 4.82% | 7.23% | 350.3951 | 387.2788
Sn, ppm | 29.2167 | 3.304 | 22.6087 | 35.8247 | 19.3047 | 39.1287 | 11.31% | 22.62% | 33.93% | 27.7558 | 30.6775
Sr, ppm | 764.6318 | 21.5656 | 721.5007 | 807.763 | 699.9351 | 829.3286 | 2.82% | 5.641% | 8.461% | 726.4002 | 802.8634
Ta, ppm | 25.6598 | 1.6503 | 22.3592 | 28.9603 | 20.709 | 30.6105 | 6.431% | 12.86% | 19.29% | 24.3768 | 26.9427
Tb, ppm | 14.3935 | 0.53793 | 13.3176 | 15.4693 | 12.7797 | 16.0073 | 3.737% | 7.475% | 11.21% | 13.6738 | 15.1131
Th, ppm | 249.0714 | 8.9058 | 231.2598 | 266.883 | 222.354 | 275.7888 | 3.576% | 7.151% | 10.73% | 236.6178 | 261.525
Ti, wt.% | 1.9075 | 0.073938 | 1.7596 | 2.0554 | 1.6857 | 2.1293 | 3.876% | 7.752% | 11.63% | 1.8122 | 2.0029
Tm, ppm | 1.2039 | 0.10316 | 0.99756 | 1.4102 | 0.8944 | 1.5133 | 8.569% | 17.14% | 25.71% | 1.1437 | 1.2641
U, ppm | 6.1476 | 0.26636 | 5.6149 | 6.6803 | 5.3485 | 6.9467 | 4.333% | 8.665% | 13% | 5.8402 | 6.455
V, ppm | 380.1715 | 18.0481 | 344.0753 | 416.2676 | 326.0272 | 434.3157 | 4.747% | 9.495% | 14.24% | 361.1629 | 399.1801
W, ppm | 3.5952 | 0.66478 | 2.2657 | 4.9248 | 1.6009 | 5.5896 | 18.49% | 36.98% | 55.47% | 3.4155 | 3.775
Y, ppm | 132.6531 | 6.4155 | 119.822 | 145.4841 | 113.4065 | 151.8997 | 4.836% | 9.673% | 14.51% | 126.0204 | 139.2857
Yb, ppm | 5.5122 | 0.36102 | 4.7901 | 6.2342 | 4.4291 | 6.5952 | 6.55% | 13.1% | 19.65% | 5.2365 | 5.7878
Zn, ppm | 268.9 | 25.9575 | 216.985 | 320.815 | 191.0276 | 346.7724 | 9.653% | 19.31% | 28.96% | 255.455 | 282.345
Zr, ppm | 593.3136 | 26.8707 | 539.5722 | 647.055 | 512.7015 | 673.9258 | 4.529% | 9.058% | 13.59% | 563.6479 | 622.9793
4-Acid Digestion
Ag, ppm | < 2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Al, wt.% | 5.7069 | 0.13173 | 5.4434 | 5.9703 | 5.3117 | 6.1021 | 2.308% | 4.616% | 6.925% | 5.4215 | 5.9922
As, ppm | 30.7827 | 1.2784 | 28.2259 | 33.3394 | 26.9475 | 34.6178 | 4.153% | 8.306% | 12.46% | 29.2435 | 32.3218
Ba, ppm | 1036.7978 | 37.6574 | 961.483 | 1112.1125 | 923.8257 | 1149.7699 | 3.632% | 7.264% | 10.9% | 984.9579 | 1088.6377
Be, ppm | 3.5673 | 0.32448 | 2.9184 | 4.2163 | 2.5939 | 4.5408 | 9.096% | 18.19% | 27.29% | 3.389 | 3.7457
Bi, ppm | 2.6935 | 0.17256 | 2.3484 | 3.0386 | 2.1758 | 3.2112 | 6.406% | 12.81% | 19.22% | 2.5588 | 2.8282
Ca, wt.% | 1.2475 | 0.08679 | 1.074 | 1.4211 | 0.98717 | 1.5079 | 6.957% | 13.91% | 20.87% | 1.1852 | 1.3099
Cd, ppm | < 0.2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ce, ppm | 4984.2902 | 180.3861 | 4623.518 | 5345.0623 | 4443.132 | 5525.4484 | 3.619% | 7.238% | 10.86% | 4735.0757 | 5233.5047
Co, ppm | 13.0237 | 0.87369 | 11.2763 | 14.7711 | 10.4026 | 15.6447 | 6.708% | 13.42% | 20.13% | 12.3725 | 13.6749
Cr, ppm | 501.5086 | 81.5901 | 338.3284 | 664.6888 | 256.7383 | 746.2788 | 16.27% | 32.54% | 48.81% | 476.4332 | 526.584
Cs, ppm | 0.45029 | 0.063532 | 0.32323 | 0.57736 | 0.25969 | 0.64089 | 14.11% | 28.22% | 42.33% | 0.42778 | 0.47281
Cu, ppm | 60.8857 | 1.942 | 57.0017 | 64.7697 | 55.0597 | 66.7117 | 3.19% | 6.379% | 9.569% | 57.8414 | 63.93
Dy, ppm | 50.5936 | 1.2121 | 48.1694 | 53.0178 | 46.9573 | 54.2299 | 2.396% | 4.792% | 7.187% | 48.0639 | 53.1233
Er, ppm | 10.8497 | 0.4193 | 10.0111 | 11.6883 | 9.5918 | 12.1076 | 3.865% | 7.729% | 11.59% | 10.3072 | 11.3922
Eu, ppm | 79.8818 | 4.4042 | 71.0733 | 88.6902 | 66.669 | 93.0945 | 5.513% | 11.03% | 16.54% | 75.8877 | 83.8758
Fe, wt.% | 33.0543 | 0.99235 | 31.0696 | 35.039 | 30.0772 | 36.0313 | 3.002% | 6.004% | 9.007% | 31.4016 | 34.707
Ga, ppm | 58.3472 | 7.0937 | 44.1597 | 72.5347 | 37.066 | 79.6284 | 12.16% | 24.32% | 36.47% | 55.4298 | 61.2646
Gd, ppm | 163.8761 | 5.5416 | 152.7929 | 174.9592 | 147.2514 | 180.5008 | 3.382% | 6.763% | 10.14% | 155.6823 | 172.0699
Hf, ppm | 7.3979 | 0.83226 | 5.7334 | 9.0624 | 4.9011 | 9.8947 | 11.25% | 22.5% | 33.75% | 7.028 | 7.7678
Ho, ppm | 6.1793 | 0.32344 | 5.5324 | 6.8262 | 5.209 | 7.1496 | 5.234% | 10.47% | 15.7% | 5.8703 | 6.4883
In, ppm | 0.74891 | 0.038885 | 0.67113 | 0.82668 | 0.63225 | 0.86556 | 5.192% | 10.38% | 15.58% | 0.71146 | 0.78635
K, wt.% | 0.11843 | 0.0053865 | 0.10766 | 0.12921 | 0.10227 | 0.13459 | 4.548% | 9.096% | 13.64% | 0.11251 | 0.12435
La, ppm | 3669.026 | 193.4809 | 3282.0642 | 4055.9878 | 3088.5834 | 4249.4686 | 5.273% | 10.55% | 15.82% | 3485.5747 | 3852.4773
Li, ppm | 10.8383 | 1.4008 | 8.0368 | 13.6399 | 6.636 | 15.0407 | 12.92% | 25.85% | 38.77% | 10.2964 | 11.3803
Lu, ppm | 0.52948 | 0.096701 | 0.33607 | 0.72288 | 0.23937 | 0.81958 | 18.26% | 36.53% | 54.79% | 0.503 | 0.55595
Mg, wt.% | 1.007 | 0.11892 | 0.76914 | 1.2448 | 0.65022 | 1.3637 | 11.81% | 23.62% | 35.43% | 0.95663 | 1.0573
Mn, ppm | 839.9573 | 63.328 | 713.3012 | 966.6133 | 649.9732 | 1029.9413 | 7.539% | 15.08% | 22.62% | 797.9594 | 881.9551
Mo, ppm | 53.4403 | 3.5067 | 46.4268 | 60.4537 | 42.9201 | 63.9605 | 6.562% | 13.12% | 19.69% | 50.7683 | 56.1123
Na, wt.% | 0.17785 | 0.012132 | 0.15359 | 0.20211 | 0.14145 | 0.21424 | 6.821% | 13.64% | 20.46% | 0.16896 | 0.18674
Nd, ppm | 2595.2838 | 125.645 | 2343.9937 | 2846.5738 | 2218.3486 | 2972.2189 | 4.841% | 9.683% | 14.52% | 2465.5196 | 2725.048
Ni, ppm | 73.8248 | 3.1147 | 67.5955 | 80.0541 | 64.4808 | 83.1687 | 4.219% | 8.438% | 12.66% | 70.1335 | 77.516
P, wt.% | 0.45372 | 0.049008 | 0.3557 | 0.55174 | 0.30669 | 0.60074 | 10.8% | 21.6% | 32.4% | 0.43103 | 0.47641
Pb, ppm | 118.5188 | 5.3267 | 107.8655 | 129.1722 | 102.5388 | 134.4989 | 4.494% | 8.989% | 13.48% | 112.5929 | 124.4448
Pr, ppm | 735.9837 | 29.8461 | 676.2915 | 795.6759 | 646.4454 | 825.522 | 4.055% | 8.111% | 12.17% | 699.1845 | 772.7829
Rb, ppm | 6.4463 | 0.60099 | 5.2443 | 7.6483 | 4.6434 | 8.2493 | 9.323% | 18.65% | 27.97% | 6.124 | 6.7686
Re, ppm | < 0.01 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, ppm | 547.9444 | 70.186 | 407.5723 | 688.3165 | 337.3863 | 758.5026 | 12.81% | 25.62% | 38.43% | 520.5472 | 575.3417
Sb, ppm | 2.172 | 0.2516 | 1.6688 | 2.6752 | 1.4173 | 2.9268 | 11.58% | 23.17% | 34.75% | 2.0634 | 2.2806
Sc, ppm | 51.4401 | 3.7349 | 43.9703 | 58.9099 | 40.2354 | 62.6448 | 7.261% | 14.52% | 21.78% | 48.8681 | 54.0121
Sm, ppm | 361.9022 | 17.3861 | 327.1299 | 396.6745 | 309.7437 | 414.0606 | 4.804% | 9.608% | 14.41% | 343.8071 | 379.9973
Sn, ppm | 23.8764 | 1.3741 | 21.1282 | 26.6246 | 19.7542 | 27.9987 | 5.755% | 11.51% | 17.27% | 22.6826 | 25.0702
Sr, ppm | 737.5233 | 58.5207 | 620.4819 | 854.5647 | 561.9612 | 913.0854 | 7.935% | 15.87% | 23.8% | 700.6471 | 774.3995
Ta, ppm | 21.8115 | 2.3785 | 17.0545 | 26.5684 | 14.6761 | 28.9469 | 10.9% | 21.81% | 32.71% | 20.7209 | 22.9021
Tb, ppm | 14.7025 | 0.62042 | 13.4616 | 15.9433 | 12.8412 | 16.5637 | 4.22% | 8.44% | 12.66% | 13.9673 | 15.4376
Te, ppm | 0.33312 | 0.061006 | 0.21111 | 0.45513 | 0.1501 | 0.51614 | 18.31% | 36.63% | 54.94% | 0.31646 | 0.34977
Th, ppm | 245.568 | 7.9044 | 229.7592 | 261.3767 | 221.8549 | 269.281 | 3.219% | 6.438% | 9.656% | 233.2896 | 257.8463
Ti, wt.% | 0.98596 | 0.10432 | 0.77732 | 1.1946 | 0.673 | 1.2989 | 10.58% | 21.16% | 31.74% | 0.93666 | 1.0353
Tl, ppm | 0.082847 | 0.010016 | 0.062814 | 0.10288 | 0.052798 | 0.1129 | 12.09% | 24.18% | 36.27% | 0.078704 | 0.086989
Tm, ppm | 1.011 | 0.089954 | 0.83113 | 1.1909 | 0.74117 | 1.2809 | 8.897% | 17.79% | 26.69% | 0.96048 | 1.0616
U, ppm | 5.7922 | 0.36416 | 5.0639 | 6.5206 | 4.6997 | 6.8847 | 6.287% | 12.57% | 18.86% | 5.5026 | 6.0818
V, ppm | 352.6418 | 26.0486 | 300.5445 | 404.739 | 274.4958 | 430.7877 | 7.387% | 14.77% | 22.16% | 335.0097 | 370.2739
W, ppm | 2.3069 | 0.33121 | 1.6444 | 2.9693 | 1.3132 | 3.3005 | 14.36% | 28.72% | 43.07% | 2.1915 | 2.4222
Y, ppm | 125.2743 | 4.7867 | 115.7009 | 134.8477 | 110.9142 | 139.6344 | 3.821% | 7.642% | 11.46% | 119.0106 | 131.538
Yb, ppm | 4.3644 | 0.35244 | 3.6595 | 5.0692 | 3.307 | 5.4217 | 8.075% | 16.15% | 24.23% | 4.1461 | 4.5826
Zn, ppm | 243.663 | 24.7633 | 194.1364 | 293.1897 | 169.373 | 317.953 | 10.16% | 20.33% | 30.49% | 231.4799 | 255.8462
Zr, ppm | 270.2672 | 30.9223 | 208.4226 | 332.1117 | 177.5004 | 363.034 | 11.44% | 22.88% | 34.32% | 256.7538 | 283.7806

## Fusion XRF
| Analytical results for Al in OREAS 462 (Indicative Value 5.87 wt.%) | [8 empty] | Indicative
Al | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 5.88 | 5.716 | 5.927 | 5.927 |  |  |  | 1
| 1 | 2 | 5.875 | 5.875 | 5.927 | 5.98 |  |  |  | 1
| 1 | 3 | 5.885 | 5.769 | 5.927 | 5.927 |  |  |  | 16
| 1 | 4 | 5.896 | 5.769 | 5.875 | 5.927 |  |  |  | 5.8669
| 1 | 5 | 5.838 | 5.716 | 5.875 | 5.927 |  |  |  | 7
| 1 | 6 | 5.901 | 5.663 | 5.875 | 5.927
| Mean |  | 5.8792 | 5.7513 | 5.901 | 5.9358
| Median |  | 5.8825 | 5.7425 | 5.901 | 5.927
| Std Dev. |  | 0.022391 | 0.072412 | 0.028482 | 0.021637
| Rel.Std.Dev. |  | 0.3809% | 1.259% | 0.4827% | 0.3645%
| PDM3 |  | 0.2099% | -1.969% | 0.5821% | 1.176%
| Z-Score (Absolute) |  | 0.26 | 3.27 | 0.26 | 1.09
| Analytical results for Ba in OREAS 462 (Indicative Value 1120 ppm) | [8 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 1160 | [5 empty] | 1
| 1 | 2 | N.A. | 1160 | [5 empty] | 2
| 1 | 3 | N.A. | 1160 | [5 empty] | 16
| 1 | 4 | N.A. | 1070 | [5 empty] | 1119.5701
| 1 | 5 | N.A. | 1070 | [5 empty] | 8
| 1 | 6 | N.A. | 1070
| Mean |  |  | 1115
| Median |  |  | 1115
| Std Dev. |  |  | 49.295
| Rel.Std.Dev. |  |  | 4.421%
| PDM3 |  |  | -0.4082%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ca in OREAS 462 (Indicative Value 1.31 wt.%) | [8 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.286 | 1.286 | 1.286 | 1.365 |  |  |  | 1
| 1 | 2 | 1.294 | 1.329 | 1.286 | 1.365 |  |  |  | 3
| 1 | 3 | 1.301 | 1.301 | 1.279 | 1.365 |  |  |  | 16
| 1 | 4 | 1.301 | 1.322 | 1.279 | 1.365 |  |  |  | 1.3135
| 1 | 5 | 1.294 | 1.322 | 1.279 | 1.372 |  |  |  | 9
| 1 | 6 | 1.294 | 1.308 | 1.286 | 1.358
| Mean |  | 1.295 | 1.3113 | 1.2825 | 1.365
| Median |  | 1.294 | 1.315 | 1.2825 | 1.365
| Std Dev. |  | 0.0055857 | 0.01612 | 0.0038341 | 0.0044272
| Rel.Std.Dev. |  | 0.4313% | 1.229% | 0.299% | 0.3243%
| PDM3 |  | -1.412% | -0.1686% | -2.364% | 3.917%
| Z-Score (Absolute) |  | 0.39 | 0.39 | 0.96 | 2.92
| Analytical results for Ce in OREAS 462 (Certified Value 5032 ppm) | [8 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 5130 | 5050 | 4870 | 5210 | 5120 | 5000 | 4880 | 1
| 1 | 2 | 5130 | 5050 | 4700 | 5050 | 5290 | 5100 | 4880 | 0
| 1 | 3 | 5130 | 5050 | 4870 | 5210 | 5040 | 5000 | 4880 | 16
| 1 | 4 | 5130 | 5050 | 4950 | 4970 | 5290 | 5000 | 4970 | 5032.3782
| 1 | 5 | 5130 | 5050 | 5290 | 4970 | 5120 | 5100 | 4880 | 7
| 1 | 6 | 5130 | 5050 | 4950 | 5130 | 5120 | 5000 | 4880
| Mean |  | 5130 | 5050 | 4938.3333 | 5090 | 5163.3333 | 5033.3333 | 4895
| Median |  | 5130 | 5050 | 4910 | 5090 | 5120 | 5000 | 4880
| Std Dev. |  | 0 | 0 | 194.9786 | 110.2724 | 102.8915 | 51.6398 | 36.7423
| Rel.Std.Dev. |  | 0% | 0% | 3.948% | 2.166% | 1.993% | 1.026% | 0.7506%
| PDM3 |  | 1.94% | 0.3502% | -1.869% | 1.145% | 2.602% | 0.01898% | -2.73%
| Z-Score (Absolute) |  | 0.67 | 0 | 1.5 | 0.34 | 0.98 | 0.11 | 1.24
| Analytical results for Cr in OREAS 462 (Indicative Value 423 ppm) | [8 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 431 | [5 empty] | 1
| 1 | 2 | N.A. | 411 | [5 empty] | 4
| 1 | 3 | N.A. | 438 | [5 empty] | 16
| 1 | 4 | N.A. | 424 | [5 empty] | 423.0813
| 1 | 5 | N.A. | 417 | [5 empty] | 10
| 1 | 6 | N.A. | 417
| Mean |  |  | 423
| Median |  |  | 420.5
| Std Dev. |  |  | 10.0598
| Rel.Std.Dev. |  |  | 2.378%
| PDM3 |  |  | -0.01922%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Dy in OREAS 462 (Indicative Value 52 ppm) | [8 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 170 | N.A. | 52 |  |  |  |  | 1
| 1 | 2 | 260 | N.A. | 52 |  |  |  |  | 5
| 1 | 3 | 170 | N.A. | 52 |  |  |  |  | 16
| 1 | 4 | 170 | N.A. | 52 |  |  |  |  | 52.2785
| 1 | 5 | 170 | N.A. | 52 |  |  |  |  | 11
| 1 | 6 | 260 | N.A. | 52
| Mean |  | 200 |  | 52
| Median |  | 170 |  | 52
| Std Dev. |  | 46.4758 |  | 0
| Rel.Std.Dev. |  | 23.24% |  | 0%
| PDM3 |  | 282.6% |  | -0.5327%
| Z-Score (Absolute) |  | 0.67 | NA | 0.67
| Analytical results for Er in OREAS 462 (Indicative Value 17.5 ppm) | [8 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 90 | N.A. | 17 |  |  |  |  | 1
| 1 | 2 | < 90 | N.A. | 17 |  |  |  |  | 6
| 1 | 3 | < 90 | N.A. | 17 |  |  |  |  | 16
| 1 | 4 | < 90 | N.A. | 17 |  |  |  |  | 17.4902
| 1 | 5 | < 90 | N.A. | 17 |  |  |  |  | 12
| 1 | 6 | < 90 | N.A. | 17
| Mean |  |  |  | 17
| Median |  |  |  | 17
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  |  | -2.802%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Eu in OREAS 462 (Indicative Value 72 ppm) | [8 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 09 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 90 | < 90 | 90 |  |  |  |  | 1
| 1 | 2 | 90 | < 90 | 90 |  |  |  |  | 7
| 1 | 3 | 90 | 90 | 90 |  |  |  |  | 16
| 1 | 4 | 90 | < 90 | 90 |  |  |  |  | 71.9694
| 1 | 5 | 90 | < 90 | 90 |  |  |  |  | 13
| 1 | 6 | 90 | < 90 | 90
| Mean |  | 90 | 90 | 90
| Median |  | 90 | 90 | 90
| Std Dev. |  | 0 |  | 0
| Rel.Std.Dev. |  | 0% |  | 0%
| PDM3 |  | 25.05% | 25.05% | 25.05%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Fe in OREAS 462 (Certified Value 34.06 wt.%) | [8 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 09 | 13 | 14 | 15 | 22 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 34.189 | 34.413 | 33.5 | 33.573 | 34.27 |  |  | 1
| 1 | 2 | 34.266 | 33.573 | 33.43 | 33.573 | 34.27 |  |  | 0
| 1 | 3 | 34.448 | 34.693 | 33.43 | 33.573 | 34.13 |  |  | 16
| 1 | 4 | 34.385 | 35.042 | 33.29 | 33.504 | 34.27 |  |  | 34.0591
| 1 | 5 | 34.308 | 34.832 | 33.29 | 33.573 | 34.27 |  |  | 8
| 1 | 6 | 34.413 | 34.832 | 33.36 | 33.504 | 34.34
| Mean |  | 34.3348 | 34.5642 | 33.3833 | 33.55 | 34.2583
| Median |  | 34.3465 | 34.7625 | 33.395 | 33.573 | 34.27
| Std Dev. |  | 0.098142 | 0.52795 | 0.084774 | 0.035631 | 0.068823
| Rel.Std.Dev. |  | 0.2858% | 1.527% | 0.2539% | 0.1062% | 0.2009%
| PDM3 |  | 0.8095% | 1.483% | -1.984% | -1.495% | 0.5849%
| Z-Score (Absolute) |  | 0.1 | 0.67 | 1.18 | 0.96 | 0
| Analytical results for Gd in OREAS 462 (Indicative Value 163 ppm) | [8 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 09 | 14 | 22 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 170 | 170 | 170 | 170 |  |  |  | 1
| 1 | 2 | 170 | 170 | 170 | 170 |  |  |  | 8
| 1 | 3 | 170 | 170 | 90 | 170 |  |  |  | 16
| 1 | 4 | 170 | 170 | 90 | 170 |  |  |  | 162.6757
| 1 | 5 | 170 | 170 | 90 | 170 |  |  |  | 14
| 1 | 6 | 170 | 170 | 170 | 170
| Mean |  | 170 | 170 | 130 | 170
| Median |  | 170 | 170 | 130 | 170
| Std Dev. |  | 0 | 0 | 43.8178 | 0
| Rel.Std.Dev. |  | 0% | 0% | 33.71% | 0%
| PDM3 |  | 4.502% | 4.502% | -20.09% | 4.502%
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Hf in OREAS 462 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Hf | Round | Replicate | Lab | [6 empty] | 1
| No | No | 13 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 9
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | 80 | [6 empty] | 15
| 1 | 6 | < 80
| Mean |  | 80
| Median |  | 80
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 462 (Indicative Value 8.73 ppm) | [8 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 2
2
| 1 | 1 | < 90 | 9 | [5 empty] | 1
| 1 | 2 | < 90 | 9 | [5 empty] | 10
| 1 | 3 | < 90 | 9 | [5 empty] | 16
| 1 | 4 | < 90 | 9 | [5 empty] | 8.7298
| 1 | 5 | < 90 | 9 | [5 empty] | 16
| 1 | 6 | < 90 | 9
| Mean |  |  | 9
| Median |  |  | 9
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 3.095%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 462 (Indicative Value 0.124 wt.%) | [8 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.116 | 0.141 | 0.116 | 0.125 |  |  |  | 1
| 1 | 2 | 0.116 | 0.108 | 0.116 | 0.116 |  |  |  | 11
| 1 | 3 | 0.116 | 0.141 | 0.116 | 0.125 |  |  |  | 16
| 1 | 4 | 0.116 | 0.141 | 0.116 | 0.125 |  |  |  | 0.12397
| 1 | 5 | 0.116 | 0.149 | 0.116 | 0.116 |  |  |  | 17
| 1 | 6 | 0.116 | 0.149 | 0.108 | 0.116
| Mean |  | 0.116 | 0.13817 | 0.11467 | 0.1205
| Median |  | 0.116 | 0.141 | 0.116 | 0.1205
| Std Dev. |  | 0 | 0.015289 | 0.003266 | 0.0049295
| Rel.Std.Dev. |  | 0% | 11.07% | 2.848% | 4.091%
| PDM3 |  | -6.428% | 11.45% | -7.504% | -2.798%
| Z-Score (Absolute) |  | 0.51 | 6.37 | 0.84 | 0.51
| Analytical results for La in OREAS 462 (Certified Value 3852 ppm) | [8 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 3840 | 3840 | 3750 | 3840 | 3920 | 4000 | 3750 | 1
| 1 | 2 | 3840 | 3840 | 3920 | 3840 | 3840 | 4000 | 3750 | 0
| 1 | 3 | 3840 | 3840 | 3920 | 3840 | 3840 | 4000 | 3840 | 16
| 1 | 4 | 3840 | 3840 | 3920 | 3750 | 3750 | 4000 | 3750 | 3851.7898
| 1 | 5 | 3750 | 3840 | 3920 | 3750 | 3840 | 4000 | 3750 | 9
| 1 | 6 | 3840 | 3840 | 3920 | 3920 | 3840 | 3900 | 3750
| Mean |  | 3825 | 3840 | 3891.6667 | 3823.3333 | 3838.3333 | 3983.3333 | 3765
| Median |  | 3840 | 3840 | 3920 | 3840 | 3840 | 4000 | 3750
| Std Dev. |  | 36.7423 | 0 | 69.4022 | 64.7045 | 53.8207 | 40.8248 | 36.7423
| Rel.Std.Dev. |  | 0.9606% | 0% | 1.783% | 1.692% | 1.402% | 1.025% | 0.9759%
| PDM3 |  | -0.6955% | -0.3061% | 1.035% | -0.7388% | -0.3494% | 3.415% | -2.253%
| Z-Score (Absolute) |  | 0.67 | 0 | 2.7 | 0.67 | 0 | 6.94 | 3.37
| Analytical results for Lu in OREAS 462 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 18 | [5 empty] | 1
| 1 | 2 | < 90 | < 18 | [5 empty] | 1
| 1 | 3 | < 90 | < 18 | [5 empty] | 16
| 1 | 4 | < 90 | < 18 | [5 empty] | < 20
| 1 | 5 | < 90 | < 18 | [5 empty] | 7
| 1 | 6 | < 90 | < 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Mg in OREAS 462 (Indicative Value 1.07 wt.%) | [8 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.079 | 1.019 | 1.11 | 1.079 |  |  |  | 1
| 1 | 2 | 1.092 | 1.043 | 1.092 | 1.061 |  |  |  | 2
| 1 | 3 | 1.086 | 1.031 | 1.073 | 1.073 |  |  |  | 16
| 1 | 4 | 1.092 | 1.043 | 1.079 | 1.079 |  |  |  | 1.0687
| 1 | 5 | 1.086 | 1.049 | 1.086 | 1.061 |  |  |  | 8
| 1 | 6 | 1.092 | 1.037 | 1.055 | 1.049
| Mean |  | 1.0878 | 1.037 | 1.0825 | 1.067
| Median |  | 1.089 | 1.04 | 1.0825 | 1.067
| Std Dev. |  | 0.0052313 | 0.010733 | 0.018534 | 0.012
| Rel.Std.Dev. |  | 0.4809% | 1.035% | 1.712% | 1.125%
| PDM3 |  | 1.792% | -2.964% | 1.293% | -0.1571%
| Z-Score (Absolute) |  | 0.84 | 2.53 | 0.51 | 0.51
| Analytical results for Mn in OREAS 462 (Indicative Value 955 ppm) | [8 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 930 | 1010 | 850 | 1010 |  |  |  | 1
| 1 | 2 | 930 | 930 | 770 | 1010 |  |  |  | 3
| 1 | 3 | 1010 | 1010 | 850 | 1010 |  |  |  | 16
| 1 | 4 | 930 | 930 | 850 | 1010 |  |  |  | 955.1838
| 1 | 5 | 1010 | 1080 | 850 | 1010 |  |  |  | 9
| 1 | 6 | 1010 | 1010 | 850 | 1010
| Mean |  | 970 | 995 | 836.6667 | 1010
| Median |  | 970 | 1010 | 850 | 1010
| Std Dev. |  | 43.8178 | 57.1839 | 32.6599 | 0
| Rel.Std.Dev. |  | 4.517% | 5.747% | 3.904% | 0%
| PDM3 |  | 1.551% | 4.168% | -12.41% | 5.739%
| Z-Score (Absolute) |  | 0.45 | 0.45 | 4.5 | 0.9
| Analytical results for Na in OREAS 462 (Indicative Value 0.191 wt.%) | [8 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.185 | 0.223 | 0.2 | 0.208 |  |  |  | 1
| 1 | 2 | 0.193 | 0.193 | 0.185 | 0.185 |  |  |  | 4
| 1 | 3 | 0.185 | 0.215 | 0.215 | 0.193 |  |  |  | 16
| 1 | 4 | 0.185 | 0.185 | 0.134 | 0.171 |  |  |  | 0.19071
| 1 | 5 | 0.178 | 0.193 | 0.185 | 0.163 |  |  |  | 10
| 1 | 6 | 0.185 | 0.185 | 0.215 | 0.215
| Mean |  | 0.18517 | 0.199 | 0.189 | 0.18917
| Median |  | 0.185 | 0.193 | 0.1925 | 0.189
| Std Dev. |  | 0.0047504 | 0.0161 | 0.0301 | 0.020341
| Rel.Std.Dev. |  | 2.565% | 8.09% | 15.93% | 10.75%
| PDM3 |  | -2.909% | 4.344% | -0.8991% | -0.8117%
| Z-Score (Absolute) |  | 1.35 | 3.6 | 0 | 0
| Analytical results for Nb in OREAS 462 (Indicative Value 1452 ppm) | [8 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | N.A. | 1470 | 1500 |  |  |  |  | 1
| 1 | 2 | N.A. | 1470 | 1500 |  |  |  |  | 5
| 1 | 3 | N.A. | 1400 | 1500 |  |  |  |  | 16
| 1 | 4 | N.A. | 1400 | 1500 |  |  |  |  | 1452.3739
| 1 | 5 | N.A. | 1400 | 1500 |  |  |  |  | 11
| 1 | 6 | N.A. | 1400 | 1400
| Mean |  |  | 1423.3333 | 1483.3333
| Median |  |  | 1400 | 1500
| Std Dev. |  |  | 36.1478 | 40.8248
| Rel.Std.Dev. |  |  | 2.54% | 2.752%
| PDM3 |  |  | -2% | 2.132%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Nd in OREAS 462 (Certified Value 2613 ppm) | [8 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 2660 | 2490 | 2400 | 2660 | 2910 | 2700 | 2581 | 1
| 1 | 2 | 2740 | 2570 | 2400 | 2570 | 2910 | 2600 | 2581 | 0
| 1 | 3 | 2660 | 2570 | 2400 | 2740 | 3000 | 2600 | 2581 | 16
| 1 | 4 | 2660 | 2570 | 2490 | 2660 | 2910 | 2600 | 2538 | 2612.8578
| 1 | 5 | 2660 | 2570 | 2570 | 2490 | 2910 | 2700 | 2563 | 10
| 1 | 6 | 2660 | 2570 | 2230 | 2660 | 3000 | 2600 | 2589
| Mean |  | 2673.3333 | 2556.6667 | 2415 | 2630 | 2940 | 2633.3333 | 2572.1667
| Median |  | 2660 | 2570 | 2400 | 2660 | 2910 | 2600 | 2581
| Std Dev. |  | 32.6599 | 32.6599 | 113.6222 | 87.178 | 46.4758 | 51.6398 | 18.8087
| Rel.Std.Dev. |  | 1.222% | 1.277% | 4.705% | 3.315% | 1.581% | 1.961% | 0.7312%
| PDM3 |  | 2.315% | -2.151% | -7.572% | 0.6561% | 12.52% | 0.7836% | -1.557%
| Z-Score (Absolute) |  | 0.51 | 0.84 | 2.53 | 0 | 3.71 | 0.05 | 0.67
| Analytical results for P in OREAS 462 (Indicative Value 0.501 wt.%) | [8 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 09 | 14 | 15 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 3
3
| 1 | 1 | 0.4944 | 0.4988 | 0.506 |  |  |  |  | 1
| 1 | 2 | 0.4944 | 0.4979 | 0.506 |  |  |  |  | 6
| 1 | 3 | 0.4944 | 0.4997 | 0.515 |  |  |  |  | 16
| 1 | 4 | 0.4953 | 0.4962 | 0.515 |  |  |  |  | 0.50086
| 1 | 5 | 0.4935 | 0.4979 | 0.506 |  |  |  |  | 12
| 1 | 6 | 0.501 | 0.4931 | 0.511
| Mean |  | 0.4955 | 0.49727 | 0.50983
| Median |  | 0.4944 | 0.4979 | 0.5085
| Std Dev. |  | 0.0027539 | 0.0023483 | 0.004446
| Rel.Std.Dev. |  | 0.5558% | 0.4722% | 0.872%
| PDM3 |  | -1.071% | -0.718% | 1.791%
| Z-Score (Absolute) |  | 0.67 | 0 | 4.86
| Analytical results for Pr in OREAS 462 (Certified Value 700 ppm) | [8 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 770 | 740 | 1110 | 500 | 850 | 700 | 740 | 1
| 1 | 2 | 770 | 740 | 1280 | 500 | 770 | 700 | 740 | 0
| 1 | 3 | 680 | 740 | 1280 | 500 | 770 | 700 | 740 | 16
| 1 | 4 | 770 | 740 | 1110 | 500 | 770 | 700 | 740 | 699.5916
| 1 | 5 | 680 | 740 | 850 | 500 | 770 | 700 | 740 | 11
| 1 | 6 | 680 | 740 | 1200 | 500 | 940 | 700 | 740
| Mean |  | 725 | 740 | 1138.3333 | 500 | 811.6667 | 700 | 740
| Median |  | 725 | 740 | 1155 | 500 | 770 | 700 | 740
| Std Dev. |  | 49.295 | 0 | 160.4265 | 0 | 70.5455 | 1.2454e-13 | 0
| Rel.Std.Dev. |  | 6.799% | 0% | 14.09% | 0% | 8.691% | 1.779e-14% | 0%
| PDM3 |  | 3.632% | 5.776% | 62.71% | -28.53% | 16.02% | 0.05838% | 5.776%
| Z-Score (Absolute) |  | 0.33 | 0 | 6.45 | 4.06 | 0.67 | 0.73 | 0
| Analytical results for Si in OREAS 462 (Indicative Value 13.33 wt.%) | [8 empty] | Indicative
Si | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 13.253 | 13.182 | 13.369 | 13.323 |  |  |  | 1
| 1 | 2 | 13.313 | 13.416 | 13.369 | 13.369 |  |  |  | 7
| 1 | 3 | 13.327 | 13.369 | 13.276 | 13.369 |  |  |  | 16
| 1 | 4 | 13.369 | 13.51 | 13.323 | 13.323 |  |  |  | 13.334
| 1 | 5 | 13.304 | 13.323 | 13.276 | 13.276 |  |  |  | 13
| 1 | 6 | 13.313 | 13.323 | 13.369 | 13.369
| Mean |  | 13.3132 | 13.3538 | 13.3303 | 13.3382
| Median |  | 13.313 | 13.346 | 13.346 | 13.346
| Std Dev. |  | 0.037451 | 0.10946 | 0.045702 | 0.037886
| Rel.Std.Dev. |  | 0.2813% | 0.8197% | 0.3428% | 0.284%
| PDM3 |  | -0.1565% | 0.1485% | -0.02776% | 0.03099%
| Z-Score (Absolute) |  | 1.21 | 1.12 | 0.22 | 0.22
| Analytical results for Sm in OREAS 462 (Certified Value < 400 ppm) | [8 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 14 | 15 | 16 | 22 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 170 | 340 | 340 | < 430 | 400 | 430 |  | 1
| 1 | 2 | 260 | 340 | 260 | 430 | 400 | 430 |  | 0
| 1 | 3 | 170 | 340 | 340 | < 430 | 400 | 430 |  | 16
| 1 | 4 | 170 | 340 | 340 | < 430 | 400 | 430 |  | < 400
| 1 | 5 | 170 | 340 | 260 | < 430 | 400 | 430 |  | 12
| 1 | 6 | 260 | 340 | 340 | < 430 | 400 | 430
| Mean |  | 200 | 340 | 313.3333 | 430 | 400 | 430
| Median |  | 170 | 340 | 340 | 430 | 400 | 430
| Std Dev. |  | 46.4758 | 0 | 41.3118 |  | 6.2269e-14 | 0
| Rel.Std.Dev. |  | 23.24% | 0% | 13.18% |  | 1.557e-14% | 0%
PDM3
| Z-Score (Absolute) |  | 1.03 | 0.11 | 0.11 | 0.91 | 0.55 | 0.8
| Analytical results for Sn in OREAS 462 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | < 80 | [5 empty] | 1
| 1 | 2 | N.A. | < 80 | [5 empty] | 8
| 1 | 3 | N.A. | < 80 | [5 empty] | 16
| 1 | 4 | N.A. | < 80 | [5 empty] | < 100
| 1 | 5 | N.A. | < 80 | [5 empty] | 14
| 1 | 6 | N.A. | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for S in OREAS 462 (Indicative Value 0.056 wt.%) | [8 empty] | Indicative
S | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.0581 | [6 empty] | 1
| 1 | 2 | 0.0549 | [6 empty] | 9
| 1 | 3 | 0.0533 | [6 empty] | 16
| 1 | 4 | 0.0557 | [6 empty] | 0.055662
| 1 | 5 | 0.0577 | [6 empty] | 15
| 1 | 6 | 0.0545
| Mean |  | 0.0557
| Median |  | 0.0553
| Std Dev. |  | 0.0018762
| Rel.Std.Dev. |  | 3.368%
| PDM3 |  | 0.06765%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 462 (Indicative Value 620 ppm) | [8 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 680 | [5 empty] | 1
| 1 | 2 | N.A. | 680 | [5 empty] | 10
| 1 | 3 | N.A. | 590 | [5 empty] | 16
| 1 | 4 | N.A. | 590 | [5 empty] | 620.1026
| 1 | 5 | N.A. | 590 | [5 empty] | 16
| 1 | 6 | N.A. | 590
| Mean |  |  | 620
| Median |  |  | 590
| Std Dev. |  |  | 46.4758
| Rel.Std.Dev. |  |  | 7.496%
| PDM3 |  |  | -0.01655%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ta in OREAS 462 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | < 80 | [5 empty] | 1
| 1 | 2 | N.A. | < 80 | [5 empty] | 11
| 1 | 3 | N.A. | < 80 | [5 empty] | 16
| 1 | 4 | N.A. | < 80 | [5 empty] | < 100
| 1 | 5 | N.A. | < 80 | [5 empty] | 17
| 1 | 6 | N.A. | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Tb in OREAS 462 (Indicative Value 17 ppm) | [8 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | 17 | [5 empty] | 1
| 1 | 2 | < 90 | 17 | [5 empty] | 1
| 1 | 3 | < 90 | 17 | [5 empty] | 16
| 1 | 4 | < 90 | 17 | [5 empty] | 17.0039
| 1 | 5 | 90 | 17 | [5 empty] | 7
| 1 | 6 | < 90 | 17
| Mean |  | 90 | 17
| Median |  | 90 | 17
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  | 429.3% | -0.023%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Th in OREAS 462 (Indicative Value 262 ppm) | [8 empty] | Indicative
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 250 | 260 | 440 | 350 | 299 | 200 |  | 1
| 1 | 2 | 230 | 260 | < 260 | 260 | 325 | 210 |  | 2
| 1 | 3 | 250 | 260 | 350 | 260 | 308 | 220 |  | 16
| 1 | 4 | 250 | 260 | < 260 | 350 | 343 | 200 |  | 261.7139
| 1 | 5 | 230 | 260 | < 260 | 260 | 343 | 200 |  | 8
| 1 | 6 | 240 | 260 | 260 | 260 | 325 | 210
| Mean |  | 241.6667 | 260 | 350 | 290 | 323.8333 | 206.6667
| Median |  | 245 | 260 | 350 | 260 | 325 | 205
| Std Dev. |  | 9.8319 | 0 | 90 | 46.4758 | 17.9155 | 8.165
| Rel.Std.Dev. |  | 4.068% | 0% | 25.71% | 16.03% | 5.532% | 3.951%
| PDM3 |  | -7.66% | -0.6549% | 33.73% | 10.81% | 23.74% | -21.03%
| Z-Score (Absolute) |  | 0.29 | 0.29 | 0.29 | 1.06 | 1.87 | 1.21
| Analytical results for Ti in OREAS 462 (Indicative Value 1.91 wt.%) | [8 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.906 | 1.852 | 1.894 | 1.978 |  |  |  | 1
| 1 | 2 | 1.906 | 1.9 | 1.87 | 1.978 |  |  |  | 3
| 1 | 3 | 1.912 | 1.894 | 1.87 | 1.984 |  |  |  | 16
| 1 | 4 | 1.918 | 1.894 | 1.87 | 1.978 |  |  |  | 1.9139
| 1 | 5 | 1.9 | 1.9 | 1.876 | 1.972 |  |  |  | 9
| 1 | 6 | 1.912 | 1.9 | 1.876 | 1.984
| Mean |  | 1.909 | 1.89 | 1.876 | 1.979
| Median |  | 1.909 | 1.897 | 1.873 | 1.978
| Std Dev. |  | 0.0062929 | 0.018847 | 0.0092952 | 0.0045166
| Rel.Std.Dev. |  | 0.3296% | 0.9972% | 0.4955% | 0.2282%
| PDM3 |  | -0.2536% | -1.246% | -1.978% | 3.404%
| Z-Score (Absolute) |  | 0.39 | 0.39 | 0.96 | 3.25
| Analytical results for Tm in OREAS 462 (Indicative Value < 10 ppm) | [8 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 9 | [5 empty] | 1
| 1 | 2 | < 90 | < 9 | [5 empty] | 4
| 1 | 3 | < 90 | < 9 | [5 empty] | 16
| 1 | 4 | < 90 | < 9 | [5 empty] | < 10
| 1 | 5 | < 90 | < 9 | [5 empty] | 10
| 1 | 6 | < 90 | < 9
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U in OREAS 462 (Indicative Value < 100 ppm) | [8 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 80 | < 250 | N.A. | 80 |  |  |  | 1
| 1 | 2 | < 80 | < 250 | N.A. | < 80 |  |  |  | 5
| 1 | 3 | < 80 | < 250 | N.A. | < 80 |  |  |  | 16
| 1 | 4 | < 80 | < 250 | N.A. | < 80 |  |  |  | < 100
| 1 | 5 | < 80 | < 250 | N.A. | 80 |  |  |  | 11
| 1 | 6 | < 80 | < 250 | N.A. | 80
| Mean |  |  |  |  | 80
| Median |  |  |  |  | 80
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for V in OREAS 462 (Indicative Value 417 ppm) | [8 empty] | Indicative
V | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 398 | [5 empty] | 1
| 1 | 2 | N.A. | 426 | [5 empty] | 6
| 1 | 3 | N.A. | 403 | [5 empty] | 16
| 1 | 4 | N.A. | 420 | [5 empty] | 417.3202
| 1 | 5 | N.A. | 415 | [5 empty] | 12
| 1 | 6 | N.A. | 443
| Mean |  |  | 417.5
| Median |  |  | 417.5
| Std Dev. |  |  | 16.2819
| Rel.Std.Dev. |  |  | 3.9%
| PDM3 |  |  | 0.04309%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for W in OREAS 462 (Indicative Value < 100 ppm) | [8 empty] | Indicative
W | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 7
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 13
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 462 (Indicative Value 88 ppm) | [8 empty] | Indicative
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 100 | 160 | 80 | 160 | 80 | 100 | 110 | 1
| 1 | 2 | 100 | 160 | 160 | 160 | < 80 | < 100 | 110 | 8
| 1 | 3 | < 100 | 160 | 80 | 160 | < 80 | 100 | 110 | 16
| 1 | 4 | 100 | 160 | 80 | 240 | 160 | < 100 | 110 | 88.4558
| 1 | 5 | < 100 | 160 | 80 | 160 | 240 | 100 | 110 | 14
| 1 | 6 | 100 | 160 | 80 | 160 | < 80 | 100 | 126
| Mean |  | 100 | 160 | 93.3333 | 173.3333 | 160 | 100 | 112.6667
| Median |  | 100 | 160 | 80 | 160 | 160 | 100 | 110
| Std Dev. |  | 0 | 0 | 32.6599 | 32.6599 | 80 | 0 | 6.532
| Rel.Std.Dev. |  | 0% | 0% | 34.99% | 18.84% | 50% | 0% | 5.798%
| PDM3 |  | 13.05% | 80.88% | 5.514% | 95.95% | 80.88% | 13.05% | 27.37%
| Z-Score (Absolute) |  | 0.27 | 2.11 | 0 | 2.53 | 0.67 | 0.27 | 0.67
| Analytical results for Yb in OREAS 462 (Indicative Value 8.78 ppm) | [8 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 2
2
| 1 | 1 | < 90 | N.A. | 9 |  |  |  |  | 1
| 1 | 2 | < 90 | N.A. | 9 |  |  |  |  | 9
| 1 | 3 | < 90 | N.A. | 9 |  |  |  |  | 16
| 1 | 4 | < 90 | N.A. | 9 |  |  |  |  | 8.7819
| 1 | 5 | < 90 | N.A. | 9 |  |  |  |  | 15
| 1 | 6 | < 90 | N.A. | 9
| Mean |  |  |  | 9
| Median |  |  |  | 9
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  |  | 2.483%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Zr in OREAS 462 (Indicative Value 531 ppm) | [8 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | N.A. | 520 | [5 empty] | 1
| 1 | 2 | N.A. | 520 | [5 empty] | 10
| 1 | 3 | N.A. | 520 | [5 empty] | 16
| 1 | 4 | N.A. | 590 | [5 empty] | 530.5498
| 1 | 5 | N.A. | 520 | [5 empty] | 16
| 1 | 6 | N.A. | 520
| Mean |  |  | 531.6667
| Median |  |  | 520
| Std Dev. |  |  | 28.5774
| Rel.Std.Dev. |  |  | 5.375%
| PDM3 |  |  | 0.2105%
| Z-Score (Absolute) |  | NA | NA
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 462 (Indicative Value 0.667 wt.%) | [8 empty] | Indicative
H2O- | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | at 1000°C | [6 empty] | 3
3
| 1 | 1 | 0.67 | [6 empty] | 1
| 1 | 2 | 0.68 | [6 empty] | 13
| 1 | 3 | 0.66 | [6 empty] | 16
| 1 | 4 | 0.67 | [6 empty] | 0.66667
| 1 | 5 | 0.65 | [6 empty] | 19
| 1 | 6 | 0.67
| Mean |  | 0.66667
| Median |  | 0.67
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 1.549%
| PDM3 |  | -4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI in OREAS 462 (Certified Value 0.771 wt.%) | [8 empty] | Certified
LOI | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 18 | wt.%
|  |  | at 1000°C | at 1000°C | at 1000°C | BF*OES | at 1000°C | BF*XRF | at 1000°C | 3
3
| 1 | 1 | 0.77 | 0.66 | 0.78 | 0.91 | 1.07 | 0.6 | 1.4773 | 1
| 1 | 2 | 0.77 | 0.69 | 0.81 | 1.07 | 1.09 | 0.632 | 1.4588 | 0
| 1 | 3 | 0.79 | 0.67 | 0.8 | 0.94 | 1.09 | 0.643 | 1.4788 | 16
| 1 | 4 | 0.79 | 0.65 | 0.79 | 0.98 | 1.1 | 0.64 | 1.5188 | 0.77067
| 1 | 5 | 0.77 | 0.68 | 0.81 | 0.92 | 1.21 | 0.626 | 1.4797 | 14
| 1 | 6 | 0.81 | 0.67 | 0.81 | 0.97 | 1.21 | 0.634 | 1.4797
| Mean |  | 0.78333 | 0.67 | 0.8 | 0.965 | 1.1283 | 0.62917 | 1.4822
| Median |  | 0.78 | 0.67 | 0.805 | 0.955 | 1.095 | 0.633 | 1.4793
| Std Dev. |  | 0.01633 | 0.014142 | 0.012649 | 0.058224 | 0.064005 | 0.015497 | 0.019657
| Rel.Std.Dev. |  | 2.085% | 2.111% | 1.581% | 6.034% | 5.673% | 2.463% | 1.326%
| PDM3 |  | 1.644% | -13.06% | 3.806% | 25.22% | 46.41% | -18.36% | 92.33%
| Z-Score (Absolute) |  | 0.07 | 0.53 | 0 | 0.67 | 1.34 | 0.67 | 2.79
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 462 (Indicative Value 9.5 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 20 | 22 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 5 | 14 | 5 | [12 empty] | 1
| 1 | 2 | < 5 | 15 | 5 | [12 empty] | 15
| 1 | 3 | < 5 | 16 | 5 | [12 empty] | 16
| 1 | 4 | < 5 | 9 | 5 | [12 empty] | 9.5
| 1 | 5 | < 5 | 16 | 5 | [12 empty] | 21
| 1 | 6 | < 5 | 14 | 5
| Mean |  |  | 14 | 5
| Median |  |  | 14.5 | 5
| Std Dev. |  |  | 2.6077 | 0
| Rel.Std.Dev. |  |  | 18.63% | 0%
| PDM3 |  |  | 47.37% | -47.37%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Al in OREAS 462 (Certified Value 5.79 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 5.837 | 5.85 | 5.73 | 5.7856 | 5.933 | 5.4103 | 5.45 | 5.88 | 5.91 | 5.89 | [5 empty] | 1
| 1 | 2 | 5.9 | 5.68 | 5.65 | 5.8616 | 5.965 | 5.429 | 5.58 | 5.94 | 5.96 | 5.98 | [5 empty] | 1
| 1 | 3 | 5.949 | 5.69 | 5.73 | 5.7855 | 5.859 | 5.4947 | 5.45 | 6.05 | 6.02 | 5.92 | [5 empty] | 16
| 1 | 4 | 5.884 | 5.77 | 5.73 | 5.9001 | 5.986 | 5.408 | 4.5 | 6.11 | 5.96 | 5.87 | [5 empty] | 5.7942
| 1 | 5 | 5.817 | 5.69 | 5.7 | 5.8847 | 5.705 | 5.4466 | 5.61 | 6.16 | 6.06 | 5.95 | [5 empty] | 16
| 1 | 6 | 5.864 | 5.69 | 5.75 | 5.7209 | 5.822 | 5.412 | 5.57 | 6.13 | 6.01 | 5.94
| Mean |  | 5.8752 | 5.7283 | 5.715 | 5.8231 | 5.8783 | 5.4334 | 5.36 | 6.045 | 5.9867 | 5.925
| Median |  | 5.874 | 5.69 | 5.73 | 5.8236 | 5.896 | 5.4205 | 5.51 | 6.08 | 5.985 | 5.93
| Std Dev. |  | 0.04714 | 0.06824 | 0.035637 | 0.069976 | 0.10553 | 0.033413 | 0.4268 | 0.11221 | 0.053541 | 0.040373
| Rel.Std.Dev. |  | 0.8024% | 1.191% | 0.6236% | 1.202% | 1.795% | 0.6149% | 7.963% | 1.856% | 0.8943% | 0.6814%
| PDM3 |  | 1.398% | -1.136% | -1.367% | 0.4988% | 1.452% | -6.226% | -7.493% | 4.329% | 3.322% | 2.258%
| Z-Score (Absolute) |  | 0.14 | 0.64 | 0.71 | 0.14 | 0.15 | 2.2 | 1.68 | 1.04 | 0.73 | 0.4
| Analytical results for As in OREAS 462 (Indicative Value 99 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 35 | < 100 | 116.0322 | 130 | 40 | 180 | [9 empty] | 1
| 1 | 2 | 32 | < 100 | 109.2276 | 110 | 40 | 190 | [9 empty] | 16
| 1 | 3 | 32 | < 100 | 103.8172 | 140 | 40 | 190 | [9 empty] | 16
| 1 | 4 | 33 | < 100 | 111.9952 | 130 | 30 | 180 | [9 empty] | 99.1665
| 1 | 5 | 30 | < 100 | 104.5476 | 120 | 40 | 170 | [9 empty] | 22
| 1 | 6 | 35 | < 100 | 112.3764 | 130 | 40 | 210
| Mean |  | 32.8333 |  | 109.666 | 126.6667 | 38.3333 | 186.6667
| Median |  | 32.5 |  | 110.6114 | 130 | 40 | 185
| Std Dev. |  | 1.9408 |  | 4.7728 | 10.328 | 4.0825 | 13.6626
| Rel.Std.Dev. |  | 5.911% |  | 4.352% | 8.154% | 10.65% | 7.319%
| PDM3 |  | -66.89% |  | 10.59% | 27.73% | -61.34% | 88.24%
| Z-Score (Absolute) |  | 0.74 | NA | 0 | 0.16 | 0.67 | 0.75
| Analytical results for B in OREAS 462 (Indicative Value 150 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*OES | PF*OES | [13 empty] | 0
0
| 1 | 1 | 152 | 140 | [13 empty] | 1
| 1 | 2 | 146 | 140 | [13 empty] | 17
| 1 | 3 | 149 | 160 | [13 empty] | 16
| 1 | 4 | 146 | 160 | [13 empty] | 150.25
| 1 | 5 | 151 | 160 | [13 empty] | 23
| 1 | 6 | 159 | 140
| Mean |  | 150.5 | 150
| Median |  | 150 | 150
| Std Dev. |  | 4.8477 | 10.9545
| Rel.Std.Dev. |  | 3.221% | 7.303%
| PDM3 |  | 0.1664% | -0.1664%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ba in OREAS 462 (Certified Value 1025 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1080 | 999 | 1090 | 1075 | 1050 | 966.6467 | 1039 | 1002 | 910.49 | 1022 | 1050 | 1061.4928 | 1000 | 1100 | 880 | 1
| 1 | 2 | 1085 | 993 | 1055 | 1125 | 1110 | 965.5429 | 986 | 995.4 | 910.01 | 1027 | 1050 | 1041.2623 | 1050 | 1050 | 880 | 2
| 1 | 3 | 1065 | 1015 | 1090 | 1095 | 1050 | 950.5901 | 988 | 988.7 | 908.645 | 1030 | 1050 | 1066.5124 | 1010 | 1070 | 880 | 16
| 1 | 4 | 1060 | 1005 | 1065 | 1080 | 1035 | 962.9472 | 990 | 1026.2 | 911.375 | 1055 | 1050 | 1032.596 | 870 | 1040 | 950 | 1024.7736
| 1 | 5 | 1095 | 943 | 1095 | 1105 | 1040 | 940.1937 | 988 | 1025.5 | 906.57 | 986 | 1050 | 1013.1455 | 1030 | 1060 | 850 | 17
| 1 | 6 | 1050 | 1010 | 1080 | 1005 | 1055 | 945.8635 | 974 | 1007.8 | 908.9975 | 992 | 1050 | 1017.5054 | 1050 | 1050 | 820
| Mean |  | 1072.5 | 994.1667 | 1079.1667 | 1080.8333 | 1056.6667 | 955.2974 | 994.1667 | 1007.6 | 909.3479 | 1018.6667 | 1050 | 1038.7524 | 1001.6667 | 1061.6667 | 876.6667
| Median |  | 1072.5 | 1002 | 1085 | 1087.5 | 1050 | 956.7687 | 988 | 1004.9 | 909.5037 | 1024.5 | 1050 | 1036.9292 | 1020 | 1055 | 880
| Std Dev. |  | 16.9558 | 26.2482 | 15.9426 | 41.2816 | 27.1416 | 11.239 | 22.7017 | 15.5166 | 1.6844 | 25.719 | 0 | 22.0866 | 67.6511 | 21.3698 | 43.2049
| Rel.Std.Dev. |  | 1.581% | 2.64% | 1.477% | 3.819% | 2.569% | 1.176% | 2.283% | 1.54% | 0.1852% | 2.525% | 0% | 2.126% | 6.754% | 2.013% | 4.928%
| PDM3 |  | 4.657% | -2.987% | 5.308% | 5.47% | 3.112% | -6.78% | -2.987% | -1.676% | -11.26% | -0.5959% | 2.462% | 1.364% | -2.255% | 3.6% | -14.45%
| Z-Score (Absolute) |  | 0.89 | 0.67 | 1.02 | 1.05 | 0.57 | 1.45 | 0.67 | 0.41 | 2.36 | 0.19 | 0.44 | 0.21 | 0 | 0.67 | 3.02
| Analytical results for Be in OREAS 462 (Certified Value 3.76 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 17 | 22 | 23 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 2
3
| 1 | 1 | 3.4988 | 4 | 3.8 | 1.9556 | 4 | 4 | < 5 | [8 empty] | 1
| 1 | 2 | 3.4531 | 3 | 3.6 | 2.0146 | 4 | 4 | < 5 | [8 empty] | 0
| 1 | 3 | 3.8602 | 4 | 3.8 | 2.0146 | 3 | 4 | < 5 | [8 empty] | 16
| 1 | 4 | 3.3662 | 4 | 3.8 | 2.2448 | 4 | 3 | < 5 | [8 empty] | 3.756
| 1 | 5 | 3.4779 | 4 | 3.5 | 1.8131 | 4 | 4 | < 5 | [8 empty] | 18
| 1 | 6 | 3.7239 | 4 | 3.8 | 1.6404 | 4 | 4 | < 5
| Mean |  | 3.5633 | 3.8333 | 3.7167 | 1.9472 | 3.8333 | 3.8333
| Median |  | 3.4883 | 4 | 3.8 | 1.9851 | 4 | 4
| Std Dev. |  | 0.18782 | 0.40825 | 0.13292 | 0.20485 | 0.40825 | 0.40825
| Rel.Std.Dev. |  | 5.271% | 10.65% | 3.576% | 10.52% | 10.65% | 10.65%
| PDM3 |  | -5.129% | 2.059% | -1.047% | -48.16% | 2.059% | 2.059%
| Z-Score (Absolute) |  | 2.45 | 0.67 | 0.67 | 21.13 | 0.67 | 0.67 | NA
| Analytical results for Bi in OREAS 462 (Certified Value 2.58 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 3.336 | 3 | 2.5 | 1.4 | 3 | 2.4 | [9 empty] | 1
| 1 | 2 | 3.2166 | 2.8 | 2.5 | 1.3 | 2 | 2.6 | [9 empty] | 0
| 1 | 3 | 2.9475 | 2.8 | 2 | 1.3 | 2 | 2.6 | [9 empty] | 16
| 1 | 4 | 2.8987 | 2.8 | 2.5 | 1.1 | 2 | 2.4 | [9 empty] | 2.5837
| 1 | 5 | 3.1361 | 2.7 | 2 | 1.4 | 2 | 2.3 | [9 empty] | 19
| 1 | 6 | 2.9762 | 2.9 | 3 | 1.3 | 2 | 2.2
| Mean |  | 3.0852 | 2.8333 | 2.4167 | 1.3 | 2.1667 | 2.4167
| Median |  | 3.0561 | 2.8 | 2.5 | 1.3 | 2 | 2.4
| Std Dev. |  | 0.17225 | 0.10328 | 0.37639 | 0.10954 | 0.40825 | 0.16021
| Rel.Std.Dev. |  | 5.583% | 3.645% | 15.57% | 8.427% | 18.84% | 6.629%
| PDM3 |  | 19.41% | 9.662% | -6.465% | -49.68% | -16.14% | -6.465%
| Z-Score (Absolute) |  | 1.35 | 0.84 | 0 | 2.26 | 0.51 | 0
| Analytical results for Ca in OREAS 462 (Certified Value 1.28 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.261 | 1.3 | 1.3 | 1.1858 | 1.258 | 1.3 | 1.2395 | 1.329 | 1.314 | 1.3 | 1.2 |  |  |  |  | 1
| 1 | 2 | 1.288 | 1.3 | 1.29 | 1.1765 | 1.251 | 1.33 | 1.2459 | 1.336 | 1.326 | 1.3 | 1.2 |  |  |  |  | 3
| 1 | 3 | 1.291 | 1.3 | 1.3 | 1.1459 | 1.244 | 1.25 | 1.2588 | 1.294 | 1.375 | 1.4 | 1.2 |  |  |  |  | 16
| 1 | 4 | 1.265 | 1.3 | 1.3 | 1.1533 | 1.265 | 1.32 | 1.2316 | 1.086 | 1.378 | 1.3 | 1.2 |  |  |  |  | 1.2785
| 1 | 5 | 1.273 | 1.3 | 1.29 | 1.1688 | 1.215 | 1.25 | 1.248 | 1.336 | 1.391 | 1.4 | 1.2 |  |  |  |  | 20
| 1 | 6 | 1.283 | 1.3 | 1.3 | 1.1312 | 1.229 | 1.38 | 1.2459 | 1.329 | 1.38 | 1.4 | 1.3
| Mean |  | 1.2768 | 1.3 | 1.2967 | 1.1603 | 1.2437 | 1.305 | 1.2449 | 1.285 | 1.3607 | 1.35 | 1.2167
| Median |  | 1.278 | 1.3 | 1.3 | 1.161 | 1.2475 | 1.31 | 1.2459 | 1.329 | 1.3765 | 1.35 | 1.2
| Std Dev. |  | 0.0124 | 0 | 0.005164 | 0.02042 | 0.018715 | 0.0501 | 0.0090604 | 0.098748 | 0.032185 | 0.054772 | 0.040825
| Rel.Std.Dev. |  | 0.9712% | 0% | 0.3983% | 1.76% | 1.505% | 3.839% | 0.7278% | 7.685% | 2.365% | 4.057% | 3.355%
| PDM3 |  | -0.1266% | 1.686% | 1.425% | -9.246% | -2.721% | 2.077% | -2.621% | 0.5122% | 6.431% | 5.597% | -4.833%
| Z-Score (Absolute) |  | 0.26 | 0.04 | 0 | 1.78 | 0.69 | 0.11 | 0.67 | 0.37 | 0.83 | 0.7 | 1.26
| Analytical results for Cd in OREAS 462 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | < 1 | < 1 | < 10 | 0.4 | [11 empty] | 1
| 1 | 2 | < 1 | < 1 | < 10 | < 0.2 | [11 empty] | 18
| 1 | 3 | < 1 | < 1 | < 10 | 0.3 | [11 empty] | 16
| 1 | 4 | < 1 | < 1 | < 10 | < 0.2 | [11 empty] | < 1
| 1 | 5 | < 1 | < 1 | < 10 | < 0.2 | [11 empty] | 24
| 1 | 6 | < 1 | < 1 | < 10 | 0.4
| Mean |  |  |  |  | 0.36667
| Median |  |  |  |  | 0.4
| Std Dev. |  |  |  |  | 0.057735
| Rel.Std.Dev. |  |  |  |  | 15.75%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Ce in OREAS 462 (Certified Value 4951 ppm) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 5110 | 4770 | 5130 | 5020 | 4960 | 4719.3137 | 5007.9 | 4784.1 | 4899.249 | 4949.5 | 5283 | 5625.1062 | 5090 | 5140 | 4840 | 1
| 1 | 2 | 5090 | 4940 | 4910 | 5200 | 5230 | 4783.9381 | 4936 | 4803 | 4909.7887 | 4709.4 | 5149 | 5635.307 | 5050 | 4890 | 5430 | 0
| 1 | 3 | 5020 | 4790 | 5020 | 5070 | 4860 | 4667.2339 | 4932.4 | 4614.8 | 4986.5817 | 4634.5 | 5083 | 5645.4309 | 5000 | 5070 | 4900 | 16
| 1 | 4 | 5100 | 4840 | 4930 | 4980 | 4880 | 4673.6686 | 4883 | 4532.3 | 4966.5701 | 4659.7 | 5144 | 5470.6194 | 4180 | 4890 | 5220 | 4951.1664
| 1 | 5 | 5140 | 4700 | 5000 | 5060 | 5080 | 4593.4468 | 4906.9 | 4793 | 4953.1324 | 4552.3 | 5122 | 5360.6097 | 5320 | 5020 | 5460 | 21
| 1 | 6 | 5030 | 5050 | 5030 | 4680 | 4890 | 4677.5474 | 4882.1 | 4508 | 4879.5373 | 4369.1 | 5109 | 5380.1109 | 5160 | 5060 | 5280
| Mean |  | 5081.6667 | 4848.3333 | 5003.3333 | 5001.6667 | 4983.3333 | 4685.8581 | 4924.7167 | 4672.5333 | 4932.4765 | 4645.75 | 5148.3333 | 5519.5307 | 4966.6667 | 5011.6667 | 5188.3333
| Median |  | 5095 | 4815 | 5010 | 5040 | 4925 | 4675.608 | 4919.65 | 4699.45 | 4931.4605 | 4647.1 | 5133 | 5547.8628 | 5070 | 5040 | 5250
| Std Dev. |  | 47.0815 | 127.0302 | 78.9092 | 174.1742 | 145.1436 | 62.9973 | 46.8609 | 137.15 | 42.2795 | 190.7379 | 70.2329 | 132.278 | 401.0819 | 101.8659 | 263.0906
| Rel.Std.Dev. |  | 0.9265% | 2.62% | 1.577% | 3.482% | 2.913% | 1.344% | 0.9515% | 2.935% | 0.8572% | 4.106% | 1.364% | 2.397% | 8.075% | 2.033% | 5.071%
| PDM3 |  | 2.636% | -2.077% | 1.054% | 1.02% | 0.6497% | -5.359% | -0.5342% | -5.628% | -0.3775% | -6.169% | 3.982% | 11.48% | 0.3131% | 1.222% | 4.79%
| Z-Score (Absolute) |  | 0.44 | 0.87 | 0 | 0.35 | 0.11 | 1.77 | 0.44 | 1.85 | 0.4 | 2 | 0.81 | 2.88 | 0.67 | 0.05 | 1.03
| Analytical results for Co in OREAS 462 (Certified Value 12.5 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 15 | 15.3298 | 13.4 | < 50 | 10.628 | 8.3 | < 10 | 20 | 12.2 | [6 empty] | 1
| 1 | 2 | 14.1 | 14.8408 | 12.5 | < 50 | 12.1219 | 9.4 | < 10 | 20 | 11.8 | [6 empty] | 0
| 1 | 3 | 14.5 | 15.4488 | 12.3 | < 50 | 10.8907 | 11.9 | < 10 | 20 | 11.5 | [6 empty] | 16
| 1 | 4 | 14.8 | 15.0547 | 12.5 | < 50 | 10.3578 | 8.3 | < 10 | < 20 | 12.7 | [6 empty] | 12.4505
| 1 | 5 | 14.3 | 15.1062 | 11.9 | < 50 | 11.2554 | 11.2 | < 10 | < 20 | 11.3 | [6 empty] | 22
| 1 | 6 | 14.5 | 15.1431 | 11.1 | < 50 | 10.2417 | 11.7 | < 10 | 20 | 10.6
| Mean |  | 14.5333 | 15.1539 | 12.2833 |  | 10.9159 | 10.1333 |  | 20 | 11.6833
| Median |  | 14.5 | 15.1247 | 12.4 |  | 10.7593 | 10.3 |  | 20 | 11.65
| Std Dev. |  | 0.3266 | 0.21365 | 0.76004 |  | 0.69559 | 1.6717 |  | 0 | 0.73052
| Rel.Std.Dev. |  | 2.247% | 1.41% | 6.188% |  | 6.372% | 16.5% |  | 0% | 6.253%
| PDM3 |  | 16.73% | 21.71% | -1.343% |  | -12.33% | -18.61% |  | 60.64% | -6.162%
| Z-Score (Absolute) |  | 0.71 | 0.9 | 0 | NA | 0.43 | 0.67 | NA | 2.42 | 0.19
| Analytical results for Cr in OREAS 462 (Certified Value 609 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 0
0
| 1 | 1 | 650 | 600 | 620 | 670 | 570 | < 1000 | 575 | 612.09 | 620 | 600 | 615.8122 | 400 | 760 | 470 | 600 | 1
| 1 | 2 | 650 | 570 | 630 | 700 | 600 | < 1000 | 556 | 609.607 | 620 | 650 | 575.3239 | 600 | 770 | 470 | 610 | 4
| 1 | 3 | 640 | 590 | 610 | 680 | 610 | < 1000 | 584 | 608.867 | 550 | 600 | 576.6907 | 600 | 760 | 470 | 620 | 16
| 1 | 4 | 640 | 580 | 630 | 670 | 610 | < 1000 | 572 | 606.457 | 620 | 600 | 579.8154 | 300 | 670 | 470 | 590 | 608.7018
| 1 | 5 | 640 | 560 | 630 | 690 | 620 | < 1000 | 563 | 609.419 | 620 | 600 | 565.1616 | 600 | 600 | 460 | 600 | 23
| 1 | 6 | 630 | 580 | 630 | 630 | 610 | < 1000 | 575 | 609.404 | 620 | 600 | 576.4675 | 600 | 780 | 470 | 620
| Mean |  | 641.6667 | 580 | 625 | 673.3333 | 603.3333 |  | 570.8333 | 609.3073 | 608.3333 | 608.3333 | 581.5452 | 516.6667 | 723.3333 | 468.3333 | 606.6667
| Median |  | 640 | 580 | 630 | 675 | 610 |  | 573.5 | 609.4115 | 620 | 600 | 576.5791 | 600 | 760 | 470 | 605
| Std Dev. |  | 7.5277 | 14.1421 | 8.3666 | 24.2212 | 17.5119 |  | 9.9079 | 1.7985 | 28.5774 | 20.4124 | 17.514 | 132.916 | 72.2957 | 4.0825 | 12.1106
| Rel.Std.Dev. |  | 1.173% | 2.438% | 1.339% | 3.597% | 2.903% |  | 1.736% | 0.2952% | 4.698% | 3.355% | 3.012% | 25.73% | 9.995% | 0.8717% | 1.996%
| PDM3 |  | 5.416% | -4.715% | 2.678% | 10.62% | -0.8819% |  | -6.221% | 0.09949% | -0.06053% | -0.06053% | -4.461% | -15.12% | 18.83% | -23.06% | -0.3343%
| Z-Score (Absolute) |  | 1.12 | 0.79 | 0.6 | 2.1 | 0.07 | NA | 1.08 | 0.12 | 0.04 | 0.09 | 0.74 | 0.17 | 4.42 | 4.26 | 0.04
| Analytical results for Cs in OREAS 462 (Certified Value 0.44 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
2
| 1 | 1 | 0.52 | 0.37 | 0.47 | 0.41 | 0.39 | 0.36159 | 0.4 | 0.5 | 0.46 | < 1 | 0.24325 | 0.3 | < 1 | 0.7 |  | 1
| 1 | 2 | 0.55 | 0.39 | 0.48 | 0.49 | 0.46 | 0.4393 | 0.4 | 0.5 | 0.41 | < 1 | 0.2621 | 0.3 | < 1 | 0.7 |  | 0
| 1 | 3 | 0.54 | 0.46 | 0.49 | 0.41 | 0.47 | 0.38553 | 0.5 | 0.5 | 0.39 | < 1 | 0.25802 | 0.3 | < 1 | 0.8 |  | 16
| 1 | 4 | 0.53 | 0.5 | 0.45 | 0.5 | 0.43 | 0.43499 | 0.4 | 0.5 | 0.37 | < 1 | 0.24305 | 0.3 | < 1 | 0.8 |  | 0.43692
| 1 | 5 | 0.5 | 0.44 | 0.48 | 0.46 | 0.47 | 0.37394 | 0.5 | 0.5 | 0.37 | < 1 | 0.22987 | 0.2 | < 1 | 0.5 |  | 24
| 1 | 6 | 0.49 | 0.46 | 0.48 | 0.41 | 0.47 | 0.40965 | 0.4 | 0.5 | 0.34 | < 1 | 0.21738 | 0.3 | < 1 | 0.5
| Mean |  | 0.52167 | 0.43667 | 0.475 | 0.44667 | 0.44833 | 0.40083 | 0.43333 | 0.5 | 0.39 |  | 0.24228 | 0.28333 |  | 0.66667
| Median |  | 0.525 | 0.45 | 0.48 | 0.435 | 0.465 | 0.39759 | 0.4 | 0.5 | 0.38 |  | 0.24315 | 0.3 |  | 0.7
| Std Dev. |  | 0.023166 | 0.048442 | 0.013784 | 0.042269 | 0.032506 | 0.032317 | 0.05164 | 0 | 0.041473 |  | 0.01683 | 0.040825 |  | 0.13663
| Rel.Std.Dev. |  | 4.441% | 11.09% | 2.902% | 9.463% | 7.251% | 8.063% | 11.92% | 0% | 10.63% |  | 6.946% | 14.41% |  | 20.49%
| PDM3 |  | 19.4% | -0.05722% | 8.716% | 2.232% | 2.613% | -8.259% | -0.8201% | 14.44% | -10.74% |  | -44.55% | -35.15% |  | 52.58%
| Z-Score (Absolute) |  | 1.17 | 0.07 | 0.56 | 0.07 | 0.27 | 0.6 | 0.12 | 0.85 | 0.75 | NA | 2.91 | 2.07 | NA | 3.28
| Analytical results for Cu in OREAS 462 (Indicative Value 72 ppm) | [16 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | 71 | < 100 | 73.3 | 56 | < 80 | 100 | 70 | [8 empty] | 1
| 1 | 2 | 85 | < 100 | 68.0285 | 303 | < 80 | 50 | 70 | [8 empty] | 19
| 1 | 3 | 74 | < 100 | 69.9768 | 24 | < 80 | 50 | 80 | [8 empty] | 16
| 1 | 4 | 73 | < 100 | 68.9063 | 49 | < 80 | 100 | 70 | [8 empty] | 72.3442
| 1 | 5 | 69 | < 100 | 69.0685 | 28 | < 80 | 100 | 70 | [8 empty] | 25
| 1 | 6 | 62 | < 100 | 72.9813 | 31 | < 80 | 50 | 70
| Mean |  | 72.3333 |  | 70.3769 | 81.8333 |  | 75 | 71.6667
| Median |  | 72 |  | 69.5227 | 40 |  | 75 | 70
| Std Dev. |  | 7.5277 |  | 2.2305 | 109.0751 |  | 27.3861 | 4.0825
| Rel.Std.Dev. |  | 10.41% |  | 3.169% | 133.3% |  | 36.51% | 5.696%
| PDM3 |  | -0.01505% |  | -2.719% | 13.12% |  | 3.671% | -0.9366%
| Z-Score (Absolute) |  | 0.35 | NA | 0.67 | 17.81 | NA | 1.74 | 0
| Analytical results for Dy in OREAS 462 (Certified Value 51 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 49.1 | 45.8 | 52.6 | 52.4 | 50.6 | 53.761 | 52.6 | 51.3 | 41.65 | 50.79 | 54 | 48.7689 | 77.1 | 55.5 | 52 | 1
| 1 | 2 | 49.7 | 51.9 | 52.5 | 55.2 | 51.6 | 56.1086 | 52.3 | 50.8 | 41.1956 | 47.36 | 54 | 48.49 | 76.6 | 52.5 | 50.4 | 5
| 1 | 3 | 47.9 | 51 | 52.7 | 52.7 | 52.1 | 52.8114 | 51.5 | 51.3 | 42.7922 | 46.64 | 52 | 48.0237 | 79.2 | 52 | 50.7 | 16
| 1 | 4 | 49.6 | 51.1 | 52 | 53.8 | 52.2 | 54.4708 | 52 | 50.6 | 42.8577 | 46.69 | 54 | 47.6006 | 79.7 | 51.5 | 56 | 51.1987
| 1 | 5 | 49.5 | 52 | 53.3 | 53.4 | 52.4 | 52.9634 | 51.2 | 51.1 | 43.027 | 45.78 | 54.5 | 46.515 | 75.4 | 52.5 | 48.4 | 25
| 1 | 6 | 48.1 | 51.9 | 52.8 | 50 | 51.6 | 54.0788 | 52.2 | 50.2 | 43.8875 | 43.9 | 52 | 46.2746 | 76.2 | 52.5 | 49.2
| Mean |  | 48.9833 | 50.6167 | 52.65 | 52.9167 | 51.75 | 54.0323 | 51.9667 | 50.8833 | 42.5683 | 46.86 | 53.4167 | 47.6121 | 77.3667 | 52.75 | 51.1167
| Median |  | 49.3 | 51.5 | 52.65 | 53.05 | 51.85 | 53.9199 | 52.1 | 50.95 | 42.825 | 46.665 | 54 | 47.8122 | 76.85 | 52.5 | 50.55
| Std Dev. |  | 0.79099 | 2.3995 | 0.42308 | 1.7348 | 0.65038 | 1.201 | 0.52409 | 0.43551 | 0.98102 | 2.2672 | 1.1143 | 1.0267 | 1.7143 | 1.4053 | 2.697
| Rel.Std.Dev. |  | 1.615% | 4.741% | 0.8036% | 3.278% | 1.257% | 2.223% | 1.009% | 0.8559% | 2.305% | 4.838% | 2.086% | 2.156% | 2.216% | 2.664% | 5.276%
| PDM3 |  | -4.327% | -1.137% | 2.835% | 3.355% | 1.077% | 5.535% | 1.5% | -0.616% | -16.86% | -8.474% | 4.332% | -7.005% | 51.11% | 3.03% | -0.1602%
| Z-Score (Absolute) |  | 1.16 | 0.07 | 0.38 | 0.49 | 0 | 0.96 | 0.09 | 0.36 | 3.85 | 1.8 | 0.7 | 1.73 | 10.73 | 0.19 | 0.67
| Analytical results for Er in OREAS 462 (Certified Value 12.2 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 11.65 | 11.3 | 12.8 | 12.6 | 12.45 | 12.7585 | 12.4 | 11.7 | 12.8486 | 11.98 | 12.7 | 10.8481 | 9.37 | 13 | 12.4 | 1
| 1 | 2 | 11.55 | 11.95 | 12.55 | 13.4 | 12.95 | 12.0879 | 12 | 11.6 | 12.2885 | 11.06 | 12.2 | 10.9171 | 9.32 | 13 | 13.1 | 6
| 1 | 3 | 11.4 | 12.05 | 12.9 | 12.45 | 12.5 | 12.9124 | 11.7 | 12.2 | 12.8277 | 10.77 | 12 | 10.7911 | 9.41 | 11 | 13 | 16
| 1 | 4 | 11.75 | 12.1 | 12.5 | 12.15 | 12.6 | 12.5227 | 12.1 | 11.9 | 12.2437 | 11.1 | 12.3 | 10.5766 | 9.49 | 12 | 12.8 | 12.2186
| 1 | 5 | 11.95 | 12 | 12.65 | 12.7 | 12.65 | 12.0995 | 11.7 | 11.8 | 13.1114 | 10.79 | 12.1 | 10.3051 | 8.98 | 12.5 | 12.9 | 26
| 1 | 6 | 11.5 | 12.3 | 12.85 | 11.75 | 12.3 | 12.1547 | 12.2 | 12.1 | 12.3264 | 10.23 | 12 | 10.2497 | 9.32 | 12 | 12.6
| Mean |  | 11.6333 | 11.95 | 12.7083 | 12.5083 | 12.575 | 12.4226 | 12.0167 | 11.8833 | 12.6077 | 10.9883 | 12.2167 | 10.6146 | 9.315 | 12.25 | 12.8
| Median |  | 11.6 | 12.025 | 12.725 | 12.525 | 12.55 | 12.3387 | 12.05 | 11.85 | 12.577 | 10.925 | 12.15 | 10.6839 | 9.345 | 12.25 | 12.85
| Std Dev. |  | 0.19664 | 0.34059 | 0.16558 | 0.55625 | 0.22079 | 0.36081 | 0.27869 | 0.23166 | 0.36706 | 0.57673 | 0.26394 | 0.28551 | 0.17604 | 0.75829 | 0.26077
| Rel.Std.Dev. |  | 1.69% | 2.85% | 1.303% | 4.447% | 1.756% | 2.904% | 2.319% | 1.949% | 2.911% | 5.249% | 2.161% | 2.69% | 1.89% | 6.19% | 2.037%
| PDM3 |  | -4.79% | -2.198% | 4.008% | 2.371% | 2.917% | 1.67% | -1.653% | -2.744% | 3.184% | -10.07% | -0.01595% | -13.13% | -23.76% | 0.2569% | 4.758%
| Z-Score (Absolute) |  | 1.1 | 0.26 | 0.93 | 0.55 | 0.67 | 0.39 | 0.38 | 0.63 | 0.74 | 2.03 | 0 | 3.01 | 5.33 | 0.06 | 1.1
| Analytical results for Eu in OREAS 462 (Certified Value 79 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 77.2 | 78.9 | 81.4 | 78.4 | 78.6 | 79.1117 | 83.9 | 75.1 | 50.3532 | 80.28 | 81 | 82.7825 | 127 | 86.5 | 78.8 | 1
| 1 | 2 | 75.8 | 78 | 79.2 | 82.6 | 81 | 77.522 | 81.3 | 75.9 | 50.6834 | 73.66 | 79.5 | 82.1978 | 124 | 81.5 | 78.7 | 7
| 1 | 3 | 74 | 77.5 | 81.6 | 79.2 | 78.2 | 78.3372 | 82.2 | 74.3 | 51.2894 | 73.25 | 77.5 | 82.091 | 128 | 79.5 | 78.4 | 16
| 1 | 4 | 76.9 | 78.7 | 78.6 | 78.9 | 77.5 | 77.4703 | 82.7 | 75 | 50.294 | 73.67 | 80 | 80.2704 | 130 | 79 | 85.2 | 78.7127
| 1 | 5 | 76.9 | 78.3 | 81.1 | 79.8 | 80.3 | 80.2642 | 81.9 | 74.5 | 51.2492 | 71.48 | 79 | 80.1621 | 123 | 79 | 77.2 | 27
| 1 | 6 | 74.7 | 78.7 | 81.6 | 73 | 77.9 | 77.6092 | 82.8 | 74.5 | 51.3166 | 68.01 | 79 | 79.8653 | 126 | 82.5 | 75
| Mean |  | 75.9167 | 78.35 | 80.5833 | 78.65 | 78.9167 | 78.3858 | 82.4667 | 74.8833 | 50.8643 | 73.3917 | 79.3333 | 81.2282 | 126.3333 | 81.3333 | 78.8833
| Median |  | 76.35 | 78.5 | 81.25 | 79.05 | 78.4 | 77.9732 | 82.45 | 74.75 | 50.9663 | 73.455 | 79.25 | 81.1807 | 126.5 | 80.5 | 78.55
| Std Dev. |  | 1.3227 | 0.5282 | 1.3303 | 3.1393 | 1.4077 | 1.1176 | 0.89144 | 0.58793 | 0.48015 | 4.0059 | 1.169 | 1.2659 | 2.582 | 2.9098 | 3.4085
| Rel.Std.Dev. |  | 1.742% | 0.6742% | 1.651% | 3.991% | 1.784% | 1.426% | 1.081% | 0.7851% | 0.944% | 5.458% | 1.474% | 1.558% | 2.044% | 3.578% | 4.321%
| PDM3 |  | -3.552% | -0.4608% | 2.377% | -0.07968% | 0.2591% | -0.4154% | 4.769% | -4.865% | -35.38% | -6.76% | 0.7885% | 3.196% | 60.5% | 3.329% | 0.2168%
| Z-Score (Absolute) |  | 0.88 | 0.17 | 0.49 | 0.25 | 0 | 0.15 | 1.04 | 1.18 | 8.18 | 1.3 | 0.12 | 0.67 | 13.83 | 0.7 | 0.38
| Analytical results for Fe in OREAS 462 (Certified Value 33.84 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 31.6954 | 34.34 | 33.3 | 34.745 | 31.755 | 35.12 | 33.4822 | 23.5 | 34.84 | 34.9 | > 25 |  |  |  |  | 1
| 1 | 2 | 32.1679 | 33.81 | 33.19 | 35.4874 | 32.811 | 35.05 | 33.6247 | 32.6 | 33.61 | 34.7 | > 25 |  |  |  |  | 0
| 1 | 3 | 32.9296 | 33.77 | 33.66 | 35.2039 | 32.266 | 34.62 | 34.1675 | 31.7 | 35.89 | 35.6 | > 25 |  |  |  |  | 16
| 1 | 4 | 32.3414 | 33.63 | 33.53 | 35.8459 | 33.14 | 35.1 | 33.5229 | 17.2 | 35.99 | 34.9 | > 25 |  |  |  |  | 33.836
| 1 | 5 | 31.84 | 33.57 | 33.4 | 35.3766 | 30.573 | 35.45 | 33.74 | 32.8 | 35.66 | 35.3 | > 25 |  |  |  |  | 28
| 1 | 6 | 31.8014 | 33.82 | 33.75 | 35.2932 | 30.454 | 34.94 | 33.6247 | 32.1 | 35.97 | 35 | > 25
| Mean |  | 32.1293 | 33.8233 | 33.4717 | 35.3253 | 31.8332 | 35.0467 | 33.6937 | 28.3167 | 35.3267 | 35.0667
| Median |  | 32.0039 | 33.79 | 33.465 | 35.3349 | 32.0105 | 35.075 | 33.6247 | 31.9 | 35.775 | 34.95
| Std Dev. |  | 0.46185 | 0.27259 | 0.21461 | 0.36123 | 1.1267 | 0.26994 | 0.249 | 6.496 | 0.94511 | 0.3266
| Rel.Std.Dev. |  | 1.437% | 0.8059% | 0.6412% | 1.023% | 3.539% | 0.7702% | 0.739% | 22.94% | 2.675% | 0.9314%
| PDM3 |  | -5.044% | -0.03732% | -1.077% | 4.402% | -5.919% | 3.578% | -0.4206% | -16.31% | 4.406% | 3.637%
| Z-Score (Absolute) |  | 0.79 | 0.03 | 0.14 | 0.76 | 0.94 | 0.63 | 0.03 | 0.71 | 0.93 | 0.64 | NA
| Analytical results for Ga in OREAS 462 (Certified Value 52 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 12 | 14 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 66.8 | 53.2 | 56.7 | 51.1 | 35.0408 | < 0.1 | 78 | 48 | [7 empty] | 1
| 1 | 2 | 61.3 | 56 | 58.6 | 54.3 | 35.4174 | 19.6 | 80 | 50 | [7 empty] | 0
| 1 | 3 | 59.9 | 55.1 | 56.8 | 57.8 | 35.7132 | 18.4 | 74 | 49 | [7 empty] | 16
| 1 | 4 | 58.5 | 54.2 | 57.2 | 57.8 | 35.2331 | 19.8 | 78 | 53 | [7 empty] | 52.2567
| 1 | 5 | 59.1 | 54.5 | 57 | 59.8 | 35.9655 | 20.2 | 78 | 48 | [7 empty] | 29
| 1 | 6 | 61.4 | 55.4 | 53.3 | 58.7 | 35.5517 | 18 | 72 | 46
| Mean |  | 61.1667 | 54.7333 | 56.6 | 56.5833 | 35.487 | 19.2 | 76.6667 | 49
| Median |  | 60.6 | 54.8 | 56.9 | 57.8 | 35.4846 | 19.6 | 78 | 48.5
| Std Dev. |  | 2.9931 | 0.98725 | 1.7584 | 3.2579 | 0.33246 | 0.94868 | 3.0111 | 2.3664
| Rel.Std.Dev. |  | 4.893% | 1.804% | 3.107% | 5.758% | 0.9369% | 4.941% | 3.928% | 4.829%
| PDM3 |  | 17.05% | 4.739% | 8.311% | 8.28% | -32.09% | -63.26% | 46.71% | -6.232%
| Z-Score (Absolute) |  | 0.53 | 0.11 | 0.11 | 0.25 | 2.47 | 4.45 | 2.57 | 0.81
| Analytical results for Gd in OREAS 462 (Certified Value 165 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 171.5 | 140.5 | 158.5 | 155.5 | 162 | 165.2 | 152.91 | 176.9691 | 169.79 | 176.5 | 165.7 | 221 | 180 | 166 |  | 1
| 1 | 2 | 171.5 | 165 | 154.5 | 163 | 167.5 | 162.3 | 153.93 | 176.9796 | 163.81 | 172 | 164.8 | 212 | 162 | 164 |  | 8
| 1 | 3 | 163.5 | 164.5 | 163.5 | 156.5 | 168.5 | 161.1 | 155.78 | 178.2007 | 161.62 | 172 | 168.3 | 209 | 172 | 164 |  | 16
| 1 | 4 | 172 | 167 | 159.5 | 156 | 169.5 | 165.2 | 155.79 | 174.3519 | 162.73 | 172 | 166.6 | 181 | 168 | 178 |  | 165.2292
| 1 | 5 | 174.5 | 166 | 157 | 159.5 | 174.5 | 165.7 | 154.08 | 175.7372 | 158.3 | 172.5 | 163.1 | 222 | 168 | 159 |  | 30
| 1 | 6 | 168.5 | 167 | 159.5 | 143 | 165.5 | 160.5 | 152.92 | 174.3429 | 151.63 | 172.5 | 167.4 | 215 | 172 | 146
| Mean |  | 170.25 | 161.6667 | 158.75 | 155.5833 | 167.9167 | 163.3333 | 154.235 | 176.0969 | 161.3133 | 172.9167 | 165.9833 | 210 | 170.3333 | 162.8333
| Median |  | 171.5 | 165.5 | 159 | 156.25 | 168 | 163.75 | 154.005 | 176.3531 | 162.175 | 172.25 | 166.15 | 213.5 | 170 | 164
| Std Dev. |  | 3.8177 | 10.4195 | 2.9958 | 6.778 | 4.1763 | 2.3088 | 1.2967 | 1.5631 | 6.0504 | 1.7725 | 1.8734 | 15.0732 | 5.9889 | 10.4003
| Rel.Std.Dev. |  | 2.242% | 6.445% | 1.887% | 4.357% | 2.487% | 1.414% | 0.8407% | 0.8877% | 3.751% | 1.025% | 1.129% | 7.178% | 3.516% | 6.387%
| PDM3 |  | 3.039% | -2.156% | -3.921% | -5.838% | 1.626% | -1.147% | -6.654% | 6.577% | -2.37% | 4.653% | 0.4564% | 27.1% | 3.089% | -1.45%
| Z-Score (Absolute) |  | 0.64 | 0.01 | 1.08 | 1.17 | 0.29 | 0.39 | 1.75 | 1.52 | 0.69 | 1.04 | 0.01 | 7.45 | 0.66 | 0.47
| Analytical results for Ge in OREAS 462 (Indicative Value 2.67 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 20 | 22 | 23 | [12 empty] | ppm
|  |  | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 2
2
| 1 | 1 | < 1 | < 20 | 3 | [12 empty] | 1
| 1 | 2 | < 1 | < 20 | 3 | [12 empty] | 21
| 1 | 3 | < 1 | < 20 | 2 | [12 empty] | 16
| 1 | 4 | < 1 | < 20 | 3 | [12 empty] | 2.6667
| 1 | 5 | < 1 | < 20 | 3 | [12 empty] | 21
| 1 | 6 | < 1 | < 20 | 2
| Mean |  |  |  | 2.6667
| Median |  |  |  | 3
| Std Dev. |  |  |  | 0.5164
| Rel.Std.Dev. |  |  |  | 19.36%
| PDM3 |  |  |  | -1.332e-13%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Hf in OREAS 462 (Certified Value 14 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 20 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 14 | 13.3 | 13.6 | 14.4 | 14.4 | 14 | 10.9 | 14.52 | 14 | 14 | 13 |  |  |  |  | 1
| 1 | 2 | 14.1 | 14.9 | 13.4 | 14.6 | 14.6 | 13.9 | 11.8 | 16.02 | 14 | 14 | 14 |  |  |  |  | 9
| 1 | 3 | 13.4 | 14.6 | 13.1 | 14.6 | 13.7 | 14.3 | 11.4 | 14.17 | 14 | 14 | 12 |  |  |  |  | 16
| 1 | 4 | 13.7 | 14.9 | 13.4 | 14.6 | 14.2 | 13.4 | 10.9 | 13.51 | 13 | 12 | 15 |  |  |  |  | 13.9638
| 1 | 5 | 14.1 | 14.4 | 14 | 14.5 | 14.7 | 14 | 11.3 | 12.57 | 14 | 15 | 14 |  |  |  |  | 31
| 1 | 6 | 13.7 | 14.7 | 13.6 | 13.4 | 13.6 | 14 | 11.6 | 11.7 | 13 | 15 | 13
| Mean |  | 13.8333 | 14.4667 | 13.5167 | 14.35 | 14.2 | 13.9333 | 11.3167 | 13.7483 | 13.6667 | 14 | 13.5
| Median |  | 13.85 | 14.65 | 13.5 | 14.55 | 14.3 | 14 | 11.35 | 13.84 | 14 | 14 | 13.5
| Std Dev. |  | 0.28048 | 0.60222 | 0.29944 | 0.47223 | 0.46043 | 0.29439 | 0.3656 | 1.5215 | 0.5164 | 1.0954 | 1.0488
| Rel.Std.Dev. |  | 2.028% | 4.163% | 2.215% | 3.291% | 3.242% | 2.113% | 3.231% | 11.07% | 3.779% | 7.825% | 7.769%
| PDM3 |  | -0.9346% | 3.601% | -3.202% | 2.765% | 1.691% | -0.2184% | -18.96% | -1.543% | -2.128% | 0.259% | -3.322%
| Z-Score (Absolute) |  | 0 | 1.85 | 0.67 | 1.5 | 0.78 | 0.21 | 5.36 | 0.18 | 0.35 | 0.35 | 0.71
| Analytical results for Ho in OREAS 462 (Certified Value 6.5 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 6.13 | 6.03 | 6.76 | 6.54 | 6.5 | 6.8377 | 7 | 6.3 | 4.2517 | 6.35 | 6.9 | 6.0111 | 7.34 | 7.6 | 6.03 | 1
| 1 | 2 | 6.1 | 6.49 | 6.76 | 6.97 | 6.68 | 7.0659 | 6.8 | 6.3 | 4.4626 | 5.89 | 6.7 | 5.978 | 6.98 | 6.8 | 5.74 | 10
| 1 | 3 | 5.86 | 6.29 | 6.79 | 6.77 | 6.55 | 6.6334 | 6.5 | 6.5 | 4.6102 | 5.95 | 6.4 | 5.9721 | 6.99 | 6.6 | 5.8 | 16
| 1 | 4 | 6.22 | 6.55 | 6.74 | 6.65 | 6.67 | 6.8925 | 6.7 | 6.3 | 4.5068 | 5.85 | 6.6 | 5.825 | 6.09 | 7 | 6.35 | 6.4951
| 1 | 5 | 6.12 | 6.59 | 6.85 | 6.92 | 6.85 | 6.5934 | 6.7 | 6.4 | 4.5563 | 5.68 | 6.6 | 5.7245 | 7.37 | 7 | 5.87 | 32
| 1 | 6 | 5.98 | 6.63 | 6.98 | 6.23 | 6.5 | 6.9258 | 6.5 | 6.3 | 4.6492 | 5.45 | 6.7 | 5.7327 | 7.16 | 6.8 | 5.25
| Mean |  | 6.0683 | 6.43 | 6.8133 | 6.68 | 6.625 | 6.8248 | 6.7 | 6.35 | 4.5061 | 5.8617 | 6.65 | 5.8739 | 6.9883 | 6.9667 | 5.84
| Median |  | 6.11 | 6.52 | 6.775 | 6.71 | 6.61 | 6.8651 | 6.7 | 6.3 | 4.5316 | 5.87 | 6.65 | 5.8985 | 7.075 | 6.9 | 5.835
| Std Dev. |  | 0.12781 | 0.22943 | 0.090259 | 0.27306 | 0.13605 | 0.18072 | 0.18974 | 0.083666 | 0.14173 | 0.29989 | 0.16432 | 0.12957 | 0.47029 | 0.34448 | 0.36244
| Rel.Std.Dev. |  | 2.106% | 3.568% | 1.325% | 4.088% | 2.054% | 2.648% | 2.832% | 1.318% | 3.145% | 5.116% | 2.471% | 2.206% | 6.73% | 4.945% | 6.206%
| PDM3 |  | -6.571% | -1.003% | 4.899% | 2.846% | 2% | 5.076% | 3.154% | -2.234% | -30.62% | -9.753% | 2.385% | -9.564% | 7.594% | 7.26% | -10.09%
| Z-Score (Absolute) |  | 1.36 | 0.28 | 0.46 | 0.13 | 0 | 0.49 | 0.18 | 0.67 | 5.2 | 1.87 | 0.06 | 1.84 | 1.33 | 0.84 | 1.92
| Analytical results for In in OREAS 462 (Indicative Value 0.78 ppm) | [16 empty] | Indicative
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.91785 | 0.9 | 0.5 | 0.5 | 1 | 0.9 | [9 empty] | 1
| 1 | 2 | 0.84243 | 0.8 | 1 | 0.5 | 0.8 | 0.9 | [9 empty] | 16
| 1 | 3 | 0.80111 | 0.8 | 0.5 | 0.5 | 0.8 | 0.8 | [9 empty] | 16
| 1 | 4 | 0.76936 | 0.8 | 0.5 | 0.3 | 0.6 | < 0.2 | [9 empty] | 0.78261
| 1 | 5 | 0.78612 | 0.8 | 0.5 | 0.5 | 0.8 | 0.8 | [9 empty] | 22
| 1 | 6 | 0.76128 | 0.8 | 0.5 | 0.5 | 0.8 | 1.1
| Mean |  | 0.81303 | 0.81667 | 0.58333 | 0.46667 | 0.8 | 0.9
| Median |  | 0.79362 | 0.8 | 0.5 | 0.5 | 0.8 | 0.9
| Std Dev. |  | 0.058827 | 0.040825 | 0.20412 | 0.08165 | 0.12649 | 0.12247
| Rel.Std.Dev. |  | 7.236% | 4.999% | 34.99% | 17.5% | 15.81% | 13.61%
| PDM3 |  | 3.887% | 4.352% | -25.46% | -40.37% | 2.223% | 15%
| Z-Score (Absolute) |  | 0.08 | 0.13 | 2.9 | 3.99 | 0.08 | 1.22
| Analytical results for K in OREAS 462 (Certified Value 0.12 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.187 | 0.15 | 0.11 | 0.12961 | 0.116 | 0.1 | 0.10335 | 0.12 | 0.12 | < 0.1 | [5 empty] | 1
| 1 | 2 | 0.187 | 0.15 | 0.1 | 0.13018 | 0.116 | 0.1 | 0.10236 | 0.13 | 0.12 | < 0.1 | [5 empty] | 11
| 1 | 3 | 0.186 | 0.13 | 0.11 | 0.13827 | 0.116 | 0.1 | 0.10144 | 0.13 | 0.13 | < 0.1 | [5 empty] | 16
| 1 | 4 | 0.195 | 0.17 | 0.11 | 0.12713 | 0.125 | 0.1 | 0.09738 | 0.11 | 0.13 | < 0.1 | [5 empty] | 0.12025
| 1 | 5 | 0.198 | 0.16 | 0.11 | 0.13791 | 0.125 | 0.1 | 0.10161 | 0.13 | 0.13 | < 0.1 | [5 empty] | 33
| 1 | 6 | 0.199 | 0.14 | 0.11 | 0.12736 | 0.116 | 0.1 | 0.10136 | 0.13 | 0.13 | < 0.1
| Mean |  | 0.192 | 0.15 | 0.10833 | 0.13174 | 0.119 | 0.1 | 0.10125 | 0.125 | 0.12667
| Median |  | 0.191 | 0.15 | 0.11 | 0.12989 | 0.116 | 0.1 | 0.10153 | 0.13 | 0.13
| Std Dev. |  | 0.006 | 0.014142 | 0.0040825 | 0.0050612 | 0.0046476 | 0 | 0.0020395 | 0.0083666 | 0.005164
| Rel.Std.Dev. |  | 3.125% | 9.428% | 3.768% | 3.842% | 3.906% | 0% | 2.014% | 6.693% | 4.077%
| PDM3 |  | 59.67% | 24.74% | -9.908% | 9.559% | -1.038% | -16.84% | -15.8% | 3.952% | 5.338%
| Z-Score (Absolute) |  | 2.71 | 1.01 | 0.67 | 0.27 | 0.24 | 1.01 | 0.96 | 0 | 0.07 | NA
| Analytical results for La in OREAS 462 (Certified Value 3794 ppm) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 3830 | 3830 | 3920 | 3720 | 3680 | 3498.77 | 3871.1 | 3749 | 3730.4582 | 3677.5 | 3943 | 4114.9264 | 3900 | 3970 | 3750 | 1
| 1 | 2 | 3790 | 3650 | 3800 | 3890 | 3830 | 3500.9257 | 3824.6 | 3737 | 3729.2934 | 3501.9 | 3844 | 4100.0412 | 3730 | 3810 | 3810 | 0
| 1 | 3 | 3750 | 3720 | 3900 | 3780 | 3710 | 3498.8951 | 3793.3 | 3767 | 3818.9428 | 3450.7 | 3795 | 4129.6647 | 3710 | 3840 | 3780 | 16
| 1 | 4 | 3800 | 3710 | 3830 | 3700 | 3760 | 3496.7231 | 3770.9 | 3761 | 3669.1812 | 3472.9 | 3838 | 4018.2168 | 3210 | 3790 | 3820 | 3794.3497
| 1 | 5 | 3840 | 3560 | 3870 | 3790 | 3860 | 3427.5568 | 3751.9 | 3744 | 3781.5513 | 3383.4 | 3819 | 3964.6243 | 3930 | 3900 | 3810 | 34
| 1 | 6 | 3750 | 3830 | 3860 | 3490 | 3750 | 3413.6784 | 3767.4 | 3782 | 3824.5545 | 3254.6 | 3823 | 3958.3724 | 3790 | 3910 | 3710
| Mean |  | 3793.3333 | 3716.6667 | 3863.3333 | 3728.3333 | 3765 | 3472.7582 | 3796.5333 | 3756.6667 | 3758.9969 | 3456.8333 | 3843.6667 | 4047.641 | 3711.6667 | 3870 | 3780
| Median |  | 3795 | 3715 | 3865 | 3750 | 3755 | 3497.7466 | 3782.1 | 3755 | 3756.0048 | 3461.8 | 3830.5 | 4059.129 | 3760 | 3870 | 3795
| Std Dev. |  | 38.2971 | 104.6263 | 44.121 | 134.3751 | 68.9202 | 40.6474 | 44.3989 | 16.5731 | 60.2721 | 139.5072 | 51.5817 | 77.107 | 261.2598 | 68.4105 | 42.8952
| Rel.Std.Dev. |  | 1.01% | 2.815% | 1.142% | 3.604% | 1.831% | 1.17% | 1.169% | 0.4412% | 1.603% | 4.036% | 1.342% | 1.905% | 7.039% | 1.768% | 1.135%
| PDM3 |  | -0.02679% | -2.047% | 1.818% | -1.74% | -0.7735% | -8.476% | 0.05755% | -0.9931% | -0.9317% | -8.895% | 1.3% | 6.675% | -2.179% | 1.994% | -0.3782%
| Z-Score (Absolute) |  | 0.28 | 1.33 | 1.76 | 0.08 | 0.32 | 6.47 | 0.35 | 0.49 | 0.44 | 6.81 | 1.34 | 5.64 | 0.67 | 1.9 | 0
| Analytical results for Li in OREAS 462 (Certified Value 11.3 ppm) | [16 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 12.4795 | 12 | 11 | 10 | 12 | < 10 | [9 empty] | 1
| 1 | 2 | 12.1193 | 12 | 11 | 20 | 12 | 10 | [9 empty] | 0
| 1 | 3 | 12.0276 | 11 | 11 | 10 | 11 | 10 | [9 empty] | 16
| 1 | 4 | 12.0292 | 12 | 13 | 20 | 11 | 10 | [9 empty] | 11.3006
| 1 | 5 | 12.1248 | 11 | 11 | 20 | 12 | 10 | [9 empty] | 35
| 1 | 6 | 12.2388 | 11 | 11 | 10 | 11 | 10
| Mean |  | 12.1699 | 11.5 | 11.3333 | 15 | 11.5 | 10
| Median |  | 12.1221 | 11.5 | 11 | 15 | 11.5 | 10
| Std Dev. |  | 0.17044 | 0.54772 | 0.8165 | 5.4772 | 0.54772 | 0
| Rel.Std.Dev. |  | 1.4% | 4.763% | 7.204% | 36.51% | 4.763% | 0%
| PDM3 |  | 7.692% | 1.764% | 0.2893% | 32.74% | 1.764% | -11.51%
| Z-Score (Absolute) |  | 1.08 | 0 | 0.27 | 5.64 | 0 | 2.42
| Analytical results for Lu in OREAS 462 (Certified Value 0.63 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 0.67 | 0.62 | 0.63 | 0.66 | 0.63 | 0.67361 | 0.8 | 0.5 | 0.63 | 1 | < 1 | 0.37 | 0.8 | 0.91 |  | 1
| 1 | 2 | 0.66 | 0.63 | 0.62 | 0.69 | 0.62 | 0.61211 | 0.6 | 0.7 | 0.59 | 0.5 | < 1 | 0.37 | 0.8 | 0.74 |  | 1
| 1 | 3 | 0.63 | 0.67 | 0.64 | 0.69 | 0.61 | 0.61236 | 0.6 | 0.7 | 0.57 | 1 | < 1 | 0.35 | 0.6 | 0.76 |  | 16
| 1 | 4 | 0.69 | 0.64 | 0.6 | 0.68 | 0.64 | 0.55987 | 0.6 | 0.7 | 0.58 | 1 | < 1 | 0.3 | 0.6 | 0.8 |  | 0.62929
| 1 | 5 | 0.66 | 0.62 | 0.67 | 0.71 | 0.64 | 0.61341 | 0.6 | 0.6 | 0.57 | 1 | < 1 | 0.38 | 0.6 | 0.73 |  | 36
| 1 | 6 | 0.64 | 0.64 | 0.67 | 0.66 | 0.64 | 0.63622 | 0.6 | 0.7 | 0.54 | 0.5 | < 1 | 0.38 | 0.6 | 0.53
| Mean |  | 0.65833 | 0.63667 | 0.63833 | 0.68167 | 0.63 | 0.61793 | 0.63333 | 0.65 | 0.58 | 0.83333 |  | 0.35833 | 0.66667 | 0.745
| Median |  | 0.66 | 0.635 | 0.635 | 0.685 | 0.635 | 0.61289 | 0.6 | 0.7 | 0.575 | 1 |  | 0.37 | 0.6 | 0.75
| Std Dev. |  | 0.02137 | 0.018619 | 0.027869 | 0.019408 | 0.012649 | 0.037127 | 0.08165 | 0.083666 | 0.029665 | 0.2582 |  | 0.030605 | 0.10328 | 0.12406
| Rel.Std.Dev. |  | 3.246% | 2.924% | 4.366% | 2.847% | 2.008% | 6.008% | 12.89% | 12.87% | 5.115% | 30.98% |  | 8.541% | 15.49% | 16.65%
| PDM3 |  | 4.615% | 1.172% | 1.437% | 8.323% | 0.1123% | -1.806% | 0.642% | 3.291% | -7.833% | 32.42% |  | -43.06% | 5.939% | 18.39%
| Z-Score (Absolute) |  | 0.4 | 0 | 0.03 | 0.83 | 0.12 | 0.34 | 0.67 | 0.25 | 1.04 | 3.62 | NA | 4.9 | 0.67 | 1.99
| Analytical results for Mg in OREAS 462 (Certified Value 1.08 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.0522 | 1.12 | 1.12 | 1.0284 | 0.965 | 1.105 | 0.95646 | 1.01 | 1.1 | 1.08 | 1.09 |  |  |  |  | 1
| 1 | 2 | 1.0626 | 1.09 | 1.11 | 1.0497 | 0.971 | 1.115 | 0.97033 | 1.06 | 1.09 | 1.07 | 1.09 |  |  |  |  | 2
| 1 | 3 | 1.0739 | 1.08 | 1.12 | 1.0477 | 0.953 | 1.095 | 0.99023 | 1.02 | 1.16 | 1.11 | 1.07 |  |  |  |  | 16
| 1 | 4 | 1.0564 | 1.1 | 1.12 | 1.05 | 0.965 | 1.11 | 0.95103 | 0.81 | 1.12 | 1.08 | 1.07 |  |  |  |  | 1.0831
| 1 | 5 | 1.0626 | 1.08 | 1.12 | 1.049 | 0.941 | 1.12 | 0.97274 | 1.05 | 1.13 | 1.11 | 1.07 |  |  |  |  | 37
| 1 | 6 | 1.0605 | 1.08 | 1.12 | 1.0292 | 0.947 | 1.11 | 0.9649 | 1.05 | 1.12 | 1.08 | 1.08
| Mean |  | 1.0613 | 1.0917 | 1.1183 | 1.0423 | 0.957 | 1.1092 | 0.96762 | 1 | 1.12 | 1.0883 | 1.0783
| Median |  | 1.0615 | 1.085 | 1.12 | 1.0483 | 0.959 | 1.11 | 0.96762 | 1.035 | 1.12 | 1.08 | 1.075
| Std Dev. |  | 0.0073501 | 0.016021 | 0.0040825 | 0.010491 | 0.011798 | 0.008612 | 0.013798 | 0.095079 | 0.024495 | 0.017224 | 0.0098319
| Rel.Std.Dev. |  | 0.6925% | 1.468% | 0.3651% | 1.007% | 1.233% | 0.7764% | 1.426% | 9.508% | 2.187% | 1.583% | 0.9118%
| PDM3 |  | -2.004% | 0.7949% | 3.257% | -3.76% | -11.64% | 2.411% | -10.66% | -7.669% | 3.411% | 0.4871% | -0.4362%
| Z-Score (Absolute) |  | 0.32 | 0.25 | 0.75 | 0.67 | 2.28 | 0.58 | 2.07 | 0.76 | 0.78 | 0.19 | 0
| Analytical results for Mn in OREAS 462 (Certified Value 933 ppm) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [5 empty] | ppm
|  |  | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | < 2000 | 1000 | 986.87 | 850 | 1000 | 827.9 | 700 | 1000 | 960 | 860 | [5 empty] | 1
| 1 | 2 | < 2000 | 900 | 990.615 | 850 | 1000 | 834.1 | 900 | 990 | 960 | 880 | [5 empty] | 3
| 1 | 3 | < 2000 | 1000 | 985.572 | 850 | 1000 | 846.5 | 900 | 1020 | 960 | 870 | [5 empty] | 16
| 1 | 4 | < 2000 | 1000 | 985.055 | 850 | 1000 | 832.6 | 500 | 1000 | 960 | 940 | [5 empty] | 932.7236
| 1 | 5 | < 2000 | 1000 | 985.924 | 850 | 1000 | 836.4 | 900 | 1010 | 960 | 850 | [5 empty] | 38
| 1 | 6 | < 2000 | 1000 | 981.377 | 850 | 1000 | 842.6 | 900 | 1010 | 960 | 830
| Mean |  |  | 983.3333 | 985.9022 | 850 | 1000 | 836.6833 | 800 | 1005 | 960 | 871.6667
| Median |  |  | 1000 | 985.748 | 850 | 1000 | 835.25 | 900 | 1005 | 960 | 865
| Std Dev. |  |  | 40.8248 | 2.9817 | 0 | 1.2454e-13 | 6.811 | 167.332 | 10.4881 | 0 | 37.6386
| Rel.Std.Dev. |  |  | 4.152% | 0.3024% | 0% | 1.245e-14% | 0.814% | 20.92% | 1.044% | 0% | 4.318%
| PDM3 |  |  | 5.426% | 5.701% | -8.869% | 7.213% | -10.3% | -14.23% | 7.749% | 2.924% | -6.546%
| Z-Score (Absolute) |  | NA | 0.35 | 0.39 | 1.62 | 0.6 | 1.85 | 0.9 | 0.67 | 0 | 1.32
| Analytical results for Mo in OREAS 462 (Certified Value 51 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
1
| 1 | 1 | 54 | 26.6923 | 53 | 55 | 48.8592 | 51 | 50 | 52 | [7 empty] | 1
| 1 | 2 | 52 | 26.8594 | 49 | 55 | 48.4505 | 51 | 50 | 55 | [7 empty] | 0
| 1 | 3 | 53 | 27.434 | 48 | 55 | 48.08 | 50 | 55 | 56 | [7 empty] | 16
| 1 | 4 | 53 | 25.1841 | 48 | 60 | 47.7262 | 42 | 50 | 48 | [7 empty] | 51.4559
| 1 | 5 | 53 | 26.8077 | 46 | 55 | 47.207 | 52 | 50 | 54 | [7 empty] | 39
| 1 | 6 | 52 | 26.9685 | 45 | 55 | 46.8231 | 51 | 55 | 52
| Mean |  | 52.8333 | 26.6577 | 48.1667 | 55.8333 | 47.8577 | 49.5 | 51.6667 | 52.8333
| Median |  | 53 | 26.8335 | 48 | 55 | 47.9031 | 51 | 50 | 53
| Std Dev. |  | 0.75277 | 0.76616 | 2.7869 | 2.0412 | 0.76353 | 3.7283 | 2.582 | 2.8577
| Rel.Std.Dev. |  | 1.425% | 2.874% | 5.786% | 3.656% | 1.595% | 7.532% | 4.997% | 5.409%
| PDM3 |  | 2.677% | -48.19% | -6.392% | 8.507% | -6.993% | -3.801% | 0.4097% | 2.677%
| Z-Score (Absolute) |  | 0.43 | 7.13 | 0.92 | 1.3 | 1 | 0.1 | 0.1 | 0.43
| Analytical results for Na in OREAS 462 (Indicative Value 0.167 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 11 | 14 | 18 | 20 | [11 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES | BF*OES | [11 empty] | 3
3
| 1 | 1 | 0.17 | 0.178 | 0.16254 | 0.84 | [11 empty] | 1
| 1 | 2 | 0.17 | 0.178 | 0.15823 | 0.85 | [11 empty] | 4
| 1 | 3 | 0.17 | 0.178 | 0.15712 | 0.82 | [11 empty] | 16
| 1 | 4 | 0.17 | 0.178 | 0.15082 | 0.64 | [11 empty] | 0.16744
| 1 | 5 | 0.17 | 0.171 | 0.15274 | 0.79 | [11 empty] | 23
| 1 | 6 | 0.17 | 0.178 | 0.15163 | 0.8
| Mean |  | 0.17 | 0.17683 | 0.15551 | 0.79
| Median |  | 0.17 | 0.178 | 0.15493 | 0.81
| Std Dev. |  | 0 | 0.0028577 | 0.0045633 | 0.076942
| Rel.Std.Dev. |  | 0% | 1.616% | 2.934% | 9.739%
| PDM3 |  | 1.529% | 5.61% | -7.122% | 371.8%
| Z-Score (Absolute) |  | 0.22 | 0.22 | 1.13 | 40.96
| Analytical results for Nb in OREAS 462 (Certified Value 1447 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 1470 | 1435 | 1425 | 1475 | 1395 | 1505 | 1417 | > 1000 | 1250 | 1080.1769 | 934 | 1430 | 1470 |  |  | 1
| 1 | 2 | 1475 | 1460 | 1420 | 1540 | 1435 | 1469 | 1356 | > 1000 | 1215 | 1189.1736 | 1120 | 1450 | 1500 |  |  | 5
| 1 | 3 | 1450 | 1470 | 1420 | 1495 | 1355 | 1468 | 1429 | > 1000 | 1155 | 1211.9792 | 1100 | 1450 | 1500 |  |  | 16
| 1 | 4 | 1455 | 1470 | 1430 | 1470 | 1365 | 1454 | 1431 | > 1000 | 1240 | 1201.8414 | 963 | 1450 | 1610 |  |  | 1447.3519
| 1 | 5 | 1465 | 1440 | 1435 | 1505 | 1420 | 1483 | 1387 | > 1000 | 1265 | 1172.0148 | 1170 | 1460 | 1440 |  |  | 40
| 1 | 6 | 1435 | 1475 | 1440 | 1375 | 1360 | 1465 | 1429 | > 1000 | 1230 | 1179.1038 | 1130 | 1470 | 1400
| Mean |  | 1458.3333 | 1458.3333 | 1428.3333 | 1476.6667 | 1388.3333 | 1474 | 1408.1667 |  | 1225.8333 | 1172.3816 | 1069.5 | 1451.6667 | 1486.6667
| Median |  | 1460 | 1465 | 1427.5 | 1485 | 1380 | 1468.5 | 1423 |  | 1235 | 1184.1387 | 1110 | 1450 | 1485
| Std Dev. |  | 14.7196 | 16.9312 | 8.165 | 55.7375 | 33.7145 | 17.7989 | 30.4527 |  | 38.6545 | 47.4655 | 96.8953 | 13.2916 | 71.4609
| Rel.Std.Dev. |  | 1.009% | 1.161% | 0.5716% | 3.775% | 2.428% | 1.208% | 2.163% |  | 3.153% | 4.049% | 9.06% | 0.9156% | 4.807%
| PDM3 |  | 0.7587% | 0.7587% | -1.314% | 2.025% | -4.078% | 1.841% | -2.707% |  | -15.31% | -19% | -26.11% | 0.2981% | 2.716%
| Z-Score (Absolute) |  | 0.38 | 0.38 | 0.24 | 1.17 | 1.06 | 0.7 | 0.65 | NA | 4.1 | 5.1 | 7.59 | 0.24 | 0.45
| Analytical results for Nd in OREAS 462 (Certified Value 2560 ppm) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 2380 | 2670 | 2670 | 2620 | 2520 | 2623.3358 | 2650.8 | 2492.7 | 2014.8941 | 2730.62 | 2709 | 2381.5544 | 2650 | 2500 | 2860 | 1
| 1 | 2 | 2390 | 2560 | 2590 | 2740 | 2610 | 2648.7631 | 2620.6 | 2533.6 | 1981.6251 | 2497.43 | 2661 | 2424.357 | 2560 | 2390 | 2870 | 0
| 1 | 3 | 2340 | 2610 | 2620 | 2650 | 2500 | 2620.0508 | 2613.7 | 2512.6 | 2095.4033 | 2454.56 | 2605 | 2434.2729 | 2520 | 2420 | 2900 | 16
| 1 | 4 | 2420 | 2600 | 2580 | 2610 | 2500 | 2626.9758 | 2588.4 | 2521 | 2018.8574 | 2460.56 | 2657 | 2354.7307 | 2180 | 2380 | 3120 | 2559.577
| 1 | 5 | 2460 | 2530 | 2650 | 2700 | 2560 | 2542.5687 | 2597 | 2493.6 | 2056.3913 | 2412.55 | 2655 | 2341.972 | 2660 | 2390 | 2820 | 41
| 1 | 6 | 2390 | 2680 | 2650 | 2440 | 2480 | 2590.879 | 2582.3 | 2498.8 | 2097.8815 | 2295.95 | 2633 | 2384.6381 | 2610 | 2420 | 2750
| Mean |  | 2396.6667 | 2608.3333 | 2626.6667 | 2626.6667 | 2528.3333 | 2608.7622 | 2608.8 | 2508.7167 | 2044.1754 | 2475.2783 | 2653.3333 | 2386.9208 | 2530 | 2416.6667 | 2886.6667
| Median |  | 2390 | 2605 | 2635 | 2635 | 2510 | 2621.6933 | 2605.35 | 2505.7 | 2037.6243 | 2457.56 | 2656 | 2383.0962 | 2585 | 2405 | 2865
| Std Dev. |  | 40.332 | 59.1326 | 36.1478 | 103.8589 | 48.3391 | 37.3444 | 25.2385 | 16.5441 | 47.0506 | 143.1289 | 34.419 | 36.6996 | 179.555 | 44.121 | 125.4857
| Rel.Std.Dev. |  | 1.683% | 2.267% | 1.376% | 3.954% | 1.912% | 1.431% | 0.9674% | 0.6595% | 2.302% | 5.782% | 1.297% | 1.538% | 7.097% | 1.826% | 4.347%
| PDM3 |  | -6.365% | 1.905% | 2.621% | 2.621% | -1.221% | 1.922% | 1.923% | -1.987% | -20.14% | -3.293% | 3.663% | -6.745% | -1.156% | -5.583% | 12.78%
| Z-Score (Absolute) |  | 1.91 | 0.08 | 0.25 | 0.25 | 0.67 | 0.08 | 0.08 | 0.86 | 5.23 | 1.65 | 0.5 | 2 | 0 | 1.72 | 2.26
| Analytical results for Ni in OREAS 462 (Certified Value 66 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 69 | 79 | < 100 | 76.7952 | 50 | 70 | 100 | 61 | [7 empty] | 1
| 1 | 2 | 45 | 88 | 200 | 69.6208 | 63 | 70 | 100 | 67 | [7 empty] | 0
| 1 | 3 | 55 | 75 | < 100 | 67.0422 | 62 | 70 | 100 | 67 | [7 empty] | 16
| 1 | 4 | 59 | 73 | < 100 | 67.8229 | 39 | 80 | 100 | 71 | [7 empty] | 66.4512
| 1 | 5 | 63 | 70 | < 100 | 67.1163 | 66 | 70 | 100 | 66 | [7 empty] | 42
| 1 | 6 | 43 | 66 | < 100 | 70.8458 | 64 | 70 | 100 | 60
| Mean |  | 55.6667 | 75.1667 | 200 | 69.8738 | 57.3333 | 71.6667 | 100 | 65.3333
| Median |  | 57 | 74 | 200 | 68.7218 | 62.5 | 70 | 100 | 66.5
| Std Dev. |  | 10.1719 | 7.679 |  | 3.7079 | 10.6145 | 4.0825 | 1.5567e-14 | 4.1312
| Rel.Std.Dev. |  | 18.27% | 10.22% |  | 5.307% | 18.51% | 5.696% | 1.557e-14% | 6.323%
| PDM3 |  | -16.23% | 13.12% | 201% | 5.151% | -13.72% | 7.849% | 50.49% | -1.682%
| Z-Score (Absolute) |  | 1.34 | 0.39 | 11.46 | 0.08 | 0.87 | 0.08 | 2.59 | 0.48
| Analytical results for P in OREAS 462 (Certified Value 0.492 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.493 | 0.5 | 0.49 | 0.48942 | 0.367 | 0.495 | 0.49267 | 0.454 | 0.48 | 0.45 | [5 empty] | 1
| 1 | 2 | 0.511 | 0.49 | 0.47 | 0.48795 | 0.358 | 0.5 | 0.49921 | 0.463 | 0.5 | 0.45 | [5 empty] | 6
| 1 | 3 | 0.505 | 0.5 | 0.49 | 0.47493 | 0.38 | 0.48 | 0.4966 | 0.445 | 0.48 | 0.45 | [5 empty] | 16
| 1 | 4 | 0.51 | 0.49 | 0.5 | 0.48521 | 0.38 | 0.485 | 0.49485 | 0.367 | 0.48 | 0.46 | [5 empty] | 0.49237
| 1 | 5 | 0.503 | 0.48 | 0.49 | 0.4934 | 0.362 | 0.49 | 0.49092 | 0.454 | 0.48 | 0.44 | [5 empty] | 43
| 1 | 6 | 0.522 | 0.49 | 0.5 | 0.48144 | 0.375 | 0.545 | 0.49878 | 0.454 | 0.5 | 0.45
| Mean |  | 0.50733 | 0.49167 | 0.49 | 0.48539 | 0.37033 | 0.49917 | 0.4955 | 0.4395 | 0.48667 | 0.45
| Median |  | 0.5075 | 0.49 | 0.49 | 0.48658 | 0.371 | 0.4925 | 0.49573 | 0.454 | 0.48 | 0.45
| Std Dev. |  | 0.0096471 | 0.0075277 | 0.010954 | 0.0065115 | 0.009395 | 0.023541 | 0.0033208 | 0.035971 | 0.010328 | 0.0063246
| Rel.Std.Dev. |  | 1.902% | 1.531% | 2.236% | 1.342% | 2.537% | 4.716% | 0.6702% | 8.184% | 2.122% | 1.405%
| PDM3 |  | 3.04% | -0.142% | -0.4805% | -1.417% | -24.78% | 1.381% | 0.6375% | -10.74% | -1.158% | -8.605%
| Z-Score (Absolute) |  | 2.44 | 0.43 | 0.21 | 0.38 | 15.17 | 0.21 | 0.92 | 4.43 | 0.21 | 4.92
| Analytical results for Pb in OREAS 462 (Certified Value 114 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 114 | 109 | 100 | 115.4888 | 62 | < 300 | 130 | 124 | [7 empty] | 1
| 1 | 2 | 112 | 116 | 200 | 114.265 | 57 | < 300 | 110 | 119 | [7 empty] | 0
| 1 | 3 | 110 | 108 | 100 | 113.2212 | 56 | < 300 | 110 | 123 | [7 empty] | 16
| 1 | 4 | 111 | 142 | 200 | 114.4837 | 49 | < 300 | 110 | 121 | [7 empty] | 114.3628
| 1 | 5 | 111 | 114 | 200 | 112.3002 | 60 | < 300 | 110 | 116 | [7 empty] | 44
| 1 | 6 | 112 | 109 | 200 | 113.9265 | 59 | < 300 | 110 | 127
| Mean |  | 111.6667 | 116.3333 | 166.6667 | 113.9476 | 57.1667 |  | 113.3333 | 121.6667
| Median |  | 111.5 | 111.5 | 200 | 114.0957 | 58 |  | 110 | 122
| Std Dev. |  | 1.3663 | 12.9718 | 51.6398 | 1.0959 | 4.535 |  | 8.165 | 3.8816
| Rel.Std.Dev. |  | 1.224% | 11.15% | 30.98% | 0.9618% | 7.933% |  | 7.204% | 3.19%
| PDM3 |  | -2.358% | 1.723% | 45.73% | -0.3631% | -50.01% |  | -0.9002% | 6.387%
| Z-Score (Absolute) |  | 0.53 | 0.67 | 16.86 | 0.19 | 17.75 | NA | 0 | 2.63
| Analytical results for Pr in OREAS 462 (Certified Value 735 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 676 | 763 | 753 | 735 | 742 | 797.566 | 757 | 763.6 | 536.695 | 735.76 | 762.5 | 747.7766 | 760 | 792 | 711 | 1
| 1 | 2 | 679 | 755 | 725 | 764 | 779 | 786.2765 | 701 | 773.7 | 540.4821 | 700.61 | 741 | 739.9704 | 727 | 740 | 723 | 0
| 1 | 3 | 666 | 760 | 741 | 745 | 722 | 802.2734 | 685 | 773.9 | 541.1448 | 688.92 | 734 | 743.4746 | 723 | 764 | 719 | 16
| 1 | 4 | 684 | 750 | 727 | 730 | 728 | 789.3186 | 714 | 769.9 | 538.2242 | 690.01 | 743.5 | 727.7575 | 612 | 746 | 772 | 735.2301
| 1 | 5 | 696 | 735 | 741 | 746 | 748 | 774.0947 | 717 | 764.9 | 538.6767 | 670.73 | 742.5 | 710.1278 | 764 | 751 | 696 | 45
| 1 | 6 | 676 | 759 | 740 | 683 | 704 | 778.6701 | 718 | 762.8 | 537.7419 | 642.77 | 735 | 709.1203 | 745 | 760 | 676
| Mean |  | 679.5 | 753.6667 | 737.8333 | 733.8333 | 737.1667 | 788.0332 | 715.3333 | 768.1333 | 538.8274 | 688.1333 | 743.0833 | 729.7045 | 721.8333 | 758.8333 | 716.1667
| Median |  | 677.5 | 757 | 740.5 | 740 | 735 | 787.7975 | 715.5 | 767.4 | 538.4504 | 689.465 | 741.75 | 733.864 | 736 | 755.5 | 715
| Std Dev. |  | 9.995 | 10.1915 | 10.3618 | 27.5058 | 25.7093 | 10.7735 | 23.9722 | 5.0361 | 1.6865 | 30.9477 | 10.2928 | 16.9254 | 56.3255 | 18.4869 | 32.3074
| Rel.Std.Dev. |  | 1.471% | 1.352% | 1.404% | 3.748% | 3.488% | 1.367% | 3.351% | 0.6556% | 0.313% | 4.497% | 1.385% | 2.319% | 7.803% | 2.436% | 4.511%
| PDM3 |  | -7.58% | 2.508% | 0.3541% | -0.19% | 0.2634% | 7.182% | -2.706% | 4.475% | -26.71% | -6.406% | 1.068% | -0.7515% | -1.822% | 3.21% | -2.593%
| Z-Score (Absolute) |  | 1.85 | 0.53 | 0.02 | 0.11 | 0 | 1.63 | 0.7 | 0.99 | 6.37 | 1.57 | 0.19 | 0.24 | 0.21 | 0.7 | 0.67
| Analytical results for Rb in OREAS 462 (Certified Value 6.58 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 6.3 | 6.4 | 6.7 | 6.3 | 6.3 | 6.6943 | 6.4 | 6.7 | 7.163 | 6.1 | 7 | 5.4 | 9 | 6.4 |  | 1
| 1 | 2 | 7 | 6.4 | 6.7 | 6.6 | 6.6 | 6.4238 | 6.3 | 6.6 | 7.2202 | 5.6 | 7 | 5.3 | 9 | 6.6 |  | 0
| 1 | 3 | 7.2 | 6.4 | 6.8 | 6.3 | 6.2 | 6.8774 | 6.5 | 6.8 | 7.1415 | 5.5 | 7 | 5.1 | 9 | 7.2 |  | 16
| 1 | 4 | 6.1 | 6.8 | 6.4 | 6.3 | 6.1 | 6.403 | 6.3 | 6.8 | 7.2494 | 5.5 | 7 | 4.6 | 8 | 7.9 |  | 6.577
| 1 | 5 | 6.3 | 5.9 | 6.6 | 6.7 | 6.3 | 6.8127 | 6.6 | 6.8 | 6.9279 | 6 | 6 | 5.5 | 9 | 6.1 |  | 46
| 1 | 6 | 5.7 | 6.3 | 6.8 | 5.7 | 6.2 | 6.832 | 5.7 | 6.7 | 7.1557 | 5.2 | 6 | 5.3 | 8 | 6.7
| Mean |  | 6.4333 | 6.3667 | 6.6667 | 6.3167 | 6.2833 | 6.6739 | 6.3 | 6.7333 | 7.1429 | 5.65 | 6.6667 | 5.2 | 8.6667 | 6.8167
| Median |  | 6.3 | 6.4 | 6.7 | 6.3 | 6.25 | 6.7535 | 6.35 | 6.75 | 7.1594 | 5.55 | 7 | 5.3 | 9 | 6.65
| Std Dev. |  | 0.56451 | 0.28752 | 0.15055 | 0.34881 | 0.17224 | 0.21072 | 0.31623 | 0.08165 | 0.11323 | 0.33912 | 0.5164 | 0.32249 | 0.5164 | 0.64317
| Rel.Std.Dev. |  | 8.775% | 4.516% | 2.258% | 5.522% | 2.741% | 3.157% | 5.019% | 1.213% | 1.585% | 6.002% | 7.746% | 6.202% | 5.958% | 9.435%
| PDM3 |  | -2.184% | -3.197% | 1.364% | -3.958% | -4.464% | 1.474% | -4.211% | 2.378% | 8.605% | -14.09% | 1.364% | -20.94% | 31.77% | 3.645%
| Z-Score (Absolute) |  | 0.31 | 0.57 | 0.57 | 0.76 | 0.88 | 0.59 | 0.36 | 0.82 | 2.53 | 3.27 | 0.57 | 4.52 | 8.12 | 0.31
| Analytical results for Re in OREAS 462 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 17 | 22 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | [12 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 1 | [12 empty] | 18
| 1 | 3 | < 0.1 | < 0.1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | [12 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | [12 empty] | 24
| 1 | 6 | < 0.1 | < 0.1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 462 (Certified Value 587 ppm) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 0
0
| 1 | 1 | 520 | 600 | 500 | 700 | 800 | [10 empty] | 1
| 1 | 2 | 590 | 600 | 600 | 500 | 600 | [10 empty] | 9
| 1 | 3 | 530 | 600 | 600 | 800 | 600 | [10 empty] | 16
| 1 | 4 | 490 | 600 | 500 | 500 | 600 | [10 empty] | 587
| 1 | 5 | 510 | 600 | 500 | 500 | 800 | [10 empty] | 47
| 1 | 6 | 570 | 600 | 600 | 500 | 600
| Mean |  | 535 | 600 | 550 | 583.3333 | 666.6667
| Median |  | 525 | 600 | 550 | 500 | 600
| Std Dev. |  | 37.8153 | 0 | 54.7723 | 132.916 | 103.2796
| Rel.Std.Dev. |  | 7.068% | 0% | 9.959% | 22.79% | 15.49%
| PDM3 |  | -8.859% | 2.215% | -6.303% | -0.6246% | 13.57%
| Z-Score (Absolute) |  | 0.98 | 0.34 | 0.67 | 0 | 1.69
| Analytical results for Sb in OREAS 462 (Certified Value 2.57 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 2.4 | 3 | 3 | 2.1634 | 3 | 2.4 | [9 empty] | 1
| 1 | 2 | 2.3 | 2.8 | 3 | 2.247 | 2 | 2.8 | [9 empty] | 0
| 1 | 3 | 2.3 | 3 | 2 | 2.2326 | 3 | 2.4 | [9 empty] | 16
| 1 | 4 | 2.5 | 3 | 3 | 2.1778 | 2 | 2.6 | [9 empty] | 2.5686
| 1 | 5 | 2.5 | 3 | 3 | 2.1195 | 3 | 2.6 | [9 empty] | 48
| 1 | 6 | 2.3 | 2.9 | 2 | 2.031 | 3 | 2.7
| Mean |  | 2.3833 | 2.95 | 2.6667 | 2.1619 | 2.6667 | 2.5833
| Median |  | 2.35 | 3 | 3 | 2.1706 | 3 | 2.6
| Std Dev. |  | 0.098319 | 0.083666 | 0.5164 | 0.079275 | 0.5164 | 0.16021
| Rel.Std.Dev. |  | 4.125% | 2.836% | 19.36% | 3.667% | 19.36% | 6.202%
| PDM3 |  | -7.214% | 14.85% | 3.816% | -15.84% | 3.816% | 0.5717%
| Z-Score (Absolute) |  | 1.15 | 1.55 | 0.2 | 2.2 | 0.2 | 0.2
| Analytical results for Sc in OREAS 462 (Indicative Value 74 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | [11 empty] | 0
0
| 1 | 1 | 55 | 100 | 60 | 110 | [11 empty] | 1
| 1 | 2 | 54 | 100 | 60 | 110 | [11 empty] | 19
| 1 | 3 | 53 | 50 | 60 | 110 | [11 empty] | 16
| 1 | 4 | 54 | 50 | 60 | 120 | [11 empty] | 74.25
| 1 | 5 | 54 | 100 | 60 | 110 | [11 empty] | 25
| 1 | 6 | 54 | 50 | 60 | 100
| Mean |  | 54 | 75 | 60 | 110
| Median |  | 54 | 75 | 60 | 110
| Std Dev. |  | 0.63246 | 27.3861 | 0 | 6.3246
| Rel.Std.Dev. |  | 1.171% | 36.51% | 0% | 5.75%
| PDM3 |  | -27.27% | 1.01% | -19.19% | 48.15%
| Z-Score (Absolute) |  | 0.87 | 0.48 | 0.48 | 2.6
| Analytical results for Se in OREAS 462 (Indicative Value < 20 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 1
1
| 1 | 1 | < 20 | < 20 | [13 empty] | 1
| 1 | 2 | < 20 | < 20 | [13 empty] | 15
| 1 | 3 | < 20 | < 20 | [13 empty] | 16
| 1 | 4 | < 20 | < 20 | [13 empty] | < 20
| 1 | 5 | < 20 | < 20 | [13 empty] | 21
| 1 | 6 | < 20 | < 20
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Si in OREAS 462 (Certified Value 13.18 wt.%) | [16 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 13.8132 | 13.04 | 13.1282 | 12.921 | 13.48 | 12.9248 | 12.949 | 12.9 | [7 empty] | 1
| 1 | 2 | 14.0703 | 12.96 | 13.2115 | 13.103 | 13.56 | 12.9648 | 13.369 | 12.8 | [7 empty] | 7
| 1 | 3 | 14.0703 | 13.07 | 13.4598 | 13.061 | 13.38 | 13.1649 | 13.603 | 13.2 | [7 empty] | 16
| 1 | 4 | 13.7 | 13.05 | 13.3101 | 13.201 | 13.62 | 12.9698 | 13.323 | 12.9 | [7 empty] | 13.1762
| 1 | 5 | 13.8033 | 12.96 | 13.1897 | 12.593 | 13.71 | 13.0299 | 13.463 | 13.1 | [7 empty] | 49
| 1 | 6 | 14.0506 | 13.08 | 13.1161 | 13.005 | 13.56 | 12.9548 | 13.65 | 12.9
| Mean |  | 13.918 | 13.0267 | 13.2359 | 12.9807 | 13.5517 | 13.0015 | 13.3928 | 12.9667
| Median |  | 13.9319 | 13.045 | 13.2006 | 13.033 | 13.56 | 12.9673 | 13.416 | 12.9
| Std Dev. |  | 0.1647 | 0.053541 | 0.12988 | 0.21185 | 0.11356 | 0.087076 | 0.25206 | 0.15055
| Rel.Std.Dev. |  | 1.183% | 0.411% | 0.9813% | 1.632% | 0.838% | 0.6697% | 1.882% | 1.161%
| PDM3 |  | 5.63% | -1.135% | 0.4531% | -1.484% | 2.85% | -1.326% | 1.644% | -1.59%
| Z-Score (Absolute) |  | 3.19 | 0.5 | 0.37 | 0.37 | 1.67 | 0.6 | 1.02 | 0.75
| Analytical results for Sm in OREAS 462 (Certified Value 369 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 365 | 350 | 380 | 376 | 353 | 368.8891 | 369.5 | 366.3 | 302.0967 | 379.5 | 376 | 379.3198 | 384 | 382 | 378 | 1
| 1 | 2 | 362 | 373 | 366 | 392 | 361 | 370.6235 | 359.9 | 373.8 | 305.4038 | 338.5 | 368.5 | 375.6499 | 370 | 361 | 373 | 0
| 1 | 3 | 355 | 379 | 373 | 384 | 353 | 374.1034 | 358.4 | 363.9 | 308.1415 | 332.2 | 364 | 374.2747 | 366 | 366 | 380 | 16
| 1 | 4 | 365 | 376 | 369 | 375 | 358 | 370.6503 | 360.4 | 368.4 | 302.8807 | 338 | 368 | 365.895 | 314 | 359 | 399 | 368.837
| 1 | 5 | 373 | 374 | 375 | 382 | 361 | 356.2454 | 357.8 | 361.8 | 305.8387 | 327.6 | 370 | 360.3585 | 387 | 362 | 365 | 50
| 1 | 6 | 363 | 379 | 377 | 351 | 353 | 371.2719 | 363.9 | 369.1 | 301.1648 | 313.4 | 366 | 363.9024 | 378 | 360 | 355
| Mean |  | 363.8333 | 371.8333 | 373.3333 | 376.6667 | 356.5 | 368.6306 | 361.65 | 367.2167 | 304.2544 | 338.2 | 368.75 | 369.9 | 366.5 | 365 | 375
| Median |  | 364 | 375 | 374 | 379 | 355.5 | 370.6369 | 360.15 | 367.35 | 304.1423 | 335.1 | 368.25 | 370.0848 | 374 | 361.5 | 375.5
| Std Dev. |  | 5.8109 | 10.9803 | 5.164 | 13.9952 | 3.9875 | 6.2998 | 4.3976 | 4.2301 | 2.6468 | 22.2226 | 4.1201 | 7.5364 | 26.9351 | 8.6718 | 14.9265
| Rel.Std.Dev. |  | 1.597% | 2.953% | 1.383% | 3.716% | 1.119% | 1.709% | 1.216% | 1.152% | 0.8699% | 6.571% | 1.117% | 2.037% | 7.349% | 2.376% | 3.98%
| PDM3 |  | -1.357% | 0.8124% | 1.219% | 2.123% | -3.345% | -0.05595% | -1.949% | -0.4393% | -17.51% | -8.306% | -0.02358% | 0.2882% | -0.6336% | -1.04% | 1.671%
| Z-Score (Absolute) |  | 0.67 | 1.06 | 0.66 | 1.13 | 1.71 | 0 | 0.98 | 0.2 | 9.05 | 5.44 | 0.02 | 0.18 | 1.18 | 0.51 | 0.22
| Analytical results for Sn in OREAS 462 (Certified Value 29.2 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS |  | 1
1
| 1 | 1 | 34 | 32 | 32 | 30 | 30 | 33 | 24 | 31.8 | 29 | 19.9117 | 58 | < 50 | 40 | 25 |  | 1
| 1 | 2 | 34 | 29 | 32 | 32 | 32 | 31 | 24 | 29.8 | 28 | 18.4594 | 29 | < 50 | 40 | 24 |  | 8
| 1 | 3 | 33 | 30 | 33 | 31 | 31 | 30 | 25 | 27.8 | 27 | 17.8093 | 28 | < 50 | 30 | 27 |  | 16
| 1 | 4 | 33 | 30 | 32 | 30 | 31 | 31 | 24 | 28.4 | 28 | 18.158 | 20 | < 50 | 40 | 26 |  | 29.2167
| 1 | 5 | 34 | 28 | 33 | 31 | 32 | 33 | 24 | 26 | 27 | 18.1803 | 31 | < 50 | 40 | 24 |  | 51
| 1 | 6 | 33 | 30 | 32 | 28 | 31 | 31 | 24 | 24.9 | 33 | 19.2015 | 31 | < 50 | 40 | 23
| Mean |  | 33.5 | 29.8333 | 32.3333 | 30.3333 | 31.1667 | 31.5 | 24.1667 | 28.1167 | 28.6667 | 18.62 | 32.8333 |  | 38.3333 | 24.8333
| Median |  | 33.5 | 30 | 32 | 30.5 | 31 | 31 | 24 | 28.1 | 28 | 18.3199 | 30 |  | 40 | 24.5
| Std Dev. |  | 0.54772 | 1.3292 | 0.5164 | 1.3663 | 0.75277 | 1.2247 | 0.40825 | 2.5063 | 2.2509 | 0.78696 | 12.9833 |  | 4.0825 | 1.472
| Rel.Std.Dev. |  | 1.635% | 4.455% | 1.597% | 4.504% | 2.415% | 3.888% | 1.689% | 8.914% | 7.852% | 4.226% | 39.54% |  | 10.65% | 5.927%
| PDM3 |  | 14.66% | 2.111% | 10.67% | 3.822% | 6.674% | 7.815% | -17.28% | -3.765% | -1.882% | -36.27% | 12.38% |  | 31.2% | -15%
| Z-Score (Absolute) |  | 1.22 | 0 | 0.83 | 0.17 | 0.44 | 0.55 | 1.88 | 0.57 | 0.67 | 3.72 | 0.67 | NA | 2.82 | 1.66
| Analytical results for Sr in OREAS 462 (Certified Value 765 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 797 | 758 | 743 | 760 | 778 | 787 | 741.3 | 740 | 800 | 912.4515 | 698 | 772 | 770 |  |  | 1
| 1 | 2 | 798 | 768 | 752 | 786 | 810 | 755 | 731.8 | 740 | 800 | 919.8911 | 674 | 779 | 770 |  |  | 10
| 1 | 3 | 782 | 766 | 756 | 769 | 753 | 744 | 759.5 | 736 | 700 | 914.7577 | 663 | 772 | 750 |  |  | 16
| 1 | 4 | 785 | 781 | 752 | 752 | 760 | 757 | 758.5 | 745 | 800 | 899.8543 | 576 | 768 | 760 |  |  | 764.6318
| 1 | 5 | 802 | 740 | 764 | 770 | 775 | 765 | 760.9 | 716 | 800 | 886.1974 | 696 | 757 | 770 |  |  | 52
| 1 | 6 | 777 | 777 | 753 | 702 | 764 | 753 | 749.3 | 734 | 800 | 894.5531 | 676 | 759 | 770
| Mean |  | 790.1667 | 765 | 753.3333 | 756.5 | 773.3333 | 760.1667 | 750.2167 | 735.1667 | 783.3333 | 904.6175 | 663.8333 | 767.8333 | 765
| Median |  | 791 | 767 | 752.5 | 764.5 | 769.5 | 756 | 753.9 | 738 | 800 | 906.1529 | 675 | 770 | 770
| Std Dev. |  | 10.1472 | 14.7241 | 6.802 | 29.0224 | 20.2353 | 14.784 | 11.7341 | 10.1275 | 40.8248 | 13.1207 | 45.0973 | 8.4242 | 8.3666
| Rel.Std.Dev. |  | 1.284% | 1.925% | 0.9029% | 3.836% | 2.617% | 1.945% | 1.564% | 1.378% | 5.212% | 1.45% | 6.793% | 1.097% | 1.094%
| PDM3 |  | 3.339% | 0.04815% | -1.478% | -1.063% | 1.138% | -0.584% | -1.885% | -3.854% | 2.446% | 18.31% | -13.18% | 0.4187% | 0.04815%
| Z-Score (Absolute) |  | 1.45 | 0 | 0.67 | 0.14 | 0.48 | 0.28 | 0.85 | 1.72 | 1.06 | 8.07 | 4.83 | 0.16 | 0
| Analytical results for Ta in OREAS 462 (Certified Value 25.7 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 1
2
| 1 | 1 | 25.4 | 22.4 | 25.6 | 26.5 | 27.9 | 26.2749 | 25.4 | 404 | 24.49 | 23 | 17.6671 | 23.9 | 27 | 25.4 |  | 1
| 1 | 2 | 25.2 | 25.5 | 25.1 | 27.2 | 28.8 | 25.6912 | 25.5 | 370 | 23.33 | 20 | 17.2627 | 28.8 | 24 | 25.7 |  | 11
| 1 | 3 | 24.7 | 25.5 | 25.3 | 26.3 | 24.8 | 25.1534 | 24.9 | 386 | 21.9 | 18 | 19.0748 | 28 | 24.5 | 26.1 |  | 16
| 1 | 4 | 24.9 | 25.7 | 25.6 | 26.4 | 28.1 | 25.6409 | 25.4 | 396 | 21.7 | 20 | 19.6024 | 18.1 | 24.5 | 28.3 |  | 25.6598
| 1 | 5 | 25.3 | 25.4 | 25.4 | 26.5 | 27.9 | 25.021 | 25 | 420 | 20.94 | 21 | 19.9741 | 30.2 | 25 | 25.6 |  | 53
| 1 | 6 | 24.7 | 26.1 | 26.2 | 25 | 27.7 | 25.8305 | 25 | 422 | 19.37 | 19 | 17.8471 | 29.2 | 25 | 24
| Mean |  | 25.0333 | 25.1 | 25.5333 | 26.3167 | 27.5333 | 25.602 | 25.2 | 399.6667 | 21.955 | 20.1667 | 18.5713 | 26.3667 | 25 | 25.85
| Median |  | 25.05 | 25.5 | 25.5 | 26.45 | 27.9 | 25.6661 | 25.2 | 400 | 21.8 | 20 | 18.4609 | 28.4 | 24.75 | 25.65
| Std Dev. |  | 0.30768 | 1.3461 | 0.37771 | 0.71949 | 1.3924 | 0.45899 | 0.26077 | 20.0566 | 1.7943 | 1.7224 | 1.126 | 4.5986 | 1.0488 | 1.3982
| Rel.Std.Dev. |  | 1.229% | 5.363% | 1.479% | 2.734% | 5.057% | 1.793% | 1.035% | 5.018% | 8.173% | 8.541% | 6.063% | 17.44% | 4.195% | 5.409%
| PDM3 |  | -2.441% | -2.181% | -0.4927% | 2.56% | 7.302% | -0.2251% | -1.792% | 1458% | -14.44% | -21.41% | -27.62% | 2.755% | -2.571% | 0.7414%
| Z-Score (Absolute) |  | 0.42 | 0.2 | 0.09 | 0.89 | 2.7 | 0.16 | 0.25 | 383.3 | 3.05 | 5.41 | 7.04 | 2.64 | 0.46 | 0.09
| Analytical results for Tb in OREAS 462 (Certified Value 14.4 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 13.95 | 12.4 | 14.4 | 13.5 | 14.3 | 15.3823 | 14.7 | 14.3 | 11.8242 | 14.71 | 14.8 | 18.4264 | 14.8 | 16.4 | 13.5 | 1
| 1 | 2 | 13.7 | 14.25 | 14.3 | 14.7 | 14.65 | 14.8684 | 13.9 | 14.5 | 11.6875 | 14.59 | 14.7 | 18.1951 | 14.2 | 15 | 14.3 | 1
| 1 | 3 | 13.05 | 13.85 | 15.2 | 13.65 | 14.85 | 15.0142 | 13.8 | 14.4 | 12.5625 | 14.33 | 14.5 | 18.0572 | 14.1 | 15.2 | 13.9 | 16
| 1 | 4 | 13.8 | 14.4 | 14.5 | 13.8 | 14.7 | 14.8179 | 14.3 | 14.5 | 12.124 | 14.57 | 14.7 | 17.8933 | 12 | 15.6 | 14.6 | 14.3935
| 1 | 5 | 13.85 | 14 | 14.65 | 13.95 | 14.95 | 15.2554 | 14.1 | 14.2 | 12.2879 | 14.19 | 14.5 | 17.4282 | 15.1 | 15.4 | 14.2 | 54
| 1 | 6 | 13.7 | 14.3 | 15.05 | 12.75 | 14.6 | 14.6029 | 13.8 | 14.3 | 12.6648 | 13.62 | 14.8 | 17.5554 | 14.4 | 14.6 | 13.1
| Mean |  | 13.675 | 13.8667 | 14.6833 | 13.725 | 14.675 | 14.9902 | 14.1 | 14.3667 | 12.1918 | 14.335 | 14.6667 | 17.9259 | 14.1 | 15.3667 | 13.9333
| Median |  | 13.75 | 14.125 | 14.575 | 13.725 | 14.675 | 14.9413 | 14 | 14.35 | 12.2059 | 14.45 | 14.7 | 17.9753 | 14.3 | 15.3 | 14.05
| Std Dev. |  | 0.32055 | 0.74677 | 0.36423 | 0.63463 | 0.22528 | 0.28955 | 0.35214 | 0.12111 | 0.39101 | 0.39828 | 0.13663 | 0.38113 | 1.0954 | 0.6121 | 0.55377
| Rel.Std.Dev. |  | 2.344% | 5.385% | 2.481% | 4.624% | 1.535% | 1.932% | 2.497% | 0.843% | 3.207% | 2.778% | 0.9315% | 2.126% | 7.769% | 3.983% | 3.974%
| PDM3 |  | -4.992% | -3.66% | 2.014% | -4.644% | 1.956% | 4.146% | -2.039% | -0.1863% | -15.3% | -0.4063% | 1.898% | 24.54% | -2.039% | 6.761% | -3.197%
| Z-Score (Absolute) |  | 1.21 | 0.44 | 0.67 | 1.37 | 0.66 | 1.33 | 0.57 | 0 | 4.63 | 0.07 | 0.64 | 7.58 | 0.33 | 1.69 | 0.92
| Analytical results for Te in OREAS 462 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 1 | < 1 | [13 empty] | 1
| 1 | 2 | < 1 | < 1 | [13 empty] | 16
| 1 | 3 | < 1 | < 1 | [13 empty] | 16
| 1 | 4 | < 1 | < 1 | [13 empty] | < 1
| 1 | 5 | < 1 | < 1 | [13 empty] | 22
| 1 | 6 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 462 (Certified Value 249 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 244 | 240 | 269 | 251 | 262 | 254.0026 | 263 | 250.7 | 193.1649 | 274.7 | 253 | 183.0448 | 186 | 240 | 241 | 1
| 1 | 2 | 245 | 246 | 256 | 260 | 274 | 252.33 | 257.7 | 253.7 | 192.2567 | 254.7 | 250 | 181.613 | 182 | 237 | 243 | 2
| 1 | 3 | 237 | 246 | 262 | 252 | 245 | 249.2033 | 257 | 251.8 | 195.6656 | 242.4 | 248.5 | 177.6598 | 179 | 238 | 245 | 16
| 1 | 4 | 245 | 249 | 251 | 252 | 249 | 252.0239 | 255.2 | 254.1 | 191.8036 | 242.6 | 249.5 | 177.5756 | 153 | 233 | 260 | 249.0714
| 1 | 5 | 245 | 244 | 263 | 254 | 254 | 248.9849 | 256.2 | 251.6 | 193.3032 | 238.6 | 250.5 | 171.3257 | 190 | 238 | 237 | 55
| 1 | 6 | 237 | 247 | 262 | 231 | 243 | 257.4366 | 254.3 | 252.5 | 192.5153 | 228 | 244.5 | 174.6746 | 183 | 239 | 229
| Mean |  | 242.1667 | 245.3333 | 260.5 | 250 | 254.5 | 252.3302 | 257.2333 | 252.4 | 193.1182 | 246.8333 | 249.3333 | 177.6489 | 178.8333 | 237.5 | 242.5
| Median |  | 244.5 | 246 | 262 | 252 | 251.5 | 252.177 | 256.6 | 252.15 | 192.8401 | 242.5 | 249.75 | 177.6177 | 182.5 | 238 | 242
| Std Dev. |  | 4.0208 | 3.0768 | 6.2209 | 9.859 | 11.7431 | 3.1602 | 3.0768 | 1.3023 | 1.3682 | 16.1152 | 2.8048 | 4.3281 | 13.1972 | 2.429 | 10.2713
| Rel.Std.Dev. |  | 1.66% | 1.254% | 2.388% | 3.944% | 4.614% | 1.252% | 1.196% | 0.516% | 0.7085% | 6.529% | 1.125% | 2.436% | 7.38% | 1.023% | 4.236%
| PDM3 |  | -2.772% | -1.501% | 4.588% | 0.3728% | 2.18% | 1.308% | 3.277% | 1.336% | -22.46% | -0.8986% | 0.1052% | -28.68% | -28.2% | -4.646% | -2.638%
| Z-Score (Absolute) |  | 0.27 | 0 | 1.31 | 0.73 | 0.79 | 0.6 | 1.02 | 0.61 | 4.49 | 0.35 | 0.34 | 5.83 | 5.28 | 0.67 | 0.24
| Analytical results for Ti in OREAS 462 (Certified Value 1.91 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.906 | 1.96 | 1.87 | 1.8779 | 1.786 | 1.95 | 1.8502 | 1.38 | 1.998 | 1.94 | 1.92 |  |  |  |  | 1
| 1 | 2 | 1.932 | 1.91 | 1.86 | 1.9053 | 1.834 | 1.96 | 1.8479 | 1.86 | 2.011 | 1.93 | 1.92 |  |  |  |  | 3
| 1 | 3 | 1.943 | 1.89 | 1.88 | 1.8818 | 1.798 | 1.93 | 1.8635 | 1.81 | 2.057 | 1.98 | 1.91 |  |  |  |  | 16
| 1 | 4 | 1.962 | 1.91 | 1.86 | 1.9104 | 1.852 | 1.97 | 1.8393 | 1.01 | 2.072 | 1.95 | 1.92 |  |  |  |  | 1.9075
| 1 | 5 | 1.921 | 1.9 | 1.86 | 1.8944 | 1.715 | 2 | 1.8497 | 1.88 | 2.085 | 1.99 | 1.93 |  |  |  |  | 56
| 1 | 6 | 1.967 | 1.9 | 1.88 | 1.8779 | 1.703 | 1.96 | 1.8445 | 1.85 | 2.081 | 1.96 | 1.93
| Mean |  | 1.9385 | 1.9117 | 1.8683 | 1.8913 | 1.7813 | 1.9617 | 1.8492 | 1.6317 | 2.0507 | 1.9583 | 1.9217
| Median |  | 1.9375 | 1.905 | 1.865 | 1.8881 | 1.792 | 1.96 | 1.8488 | 1.83 | 2.0645 | 1.955 | 1.92
| Std Dev. |  | 0.02362 | 0.024833 | 0.0098319 | 0.014285 | 0.060997 | 0.023166 | 0.008085 | 0.35863 | 0.037259 | 0.023166 | 0.0075277
| Rel.Std.Dev. |  | 1.218% | 1.299% | 0.5262% | 0.7553% | 3.424% | 1.181% | 0.4372% | 21.98% | 1.817% | 1.183% | 0.3917%
| PDM3 |  | 1.624% | 0.217% | -2.055% | -0.8509% | -6.616% | 2.838% | -3.058% | -14.46% | 7.504% | 2.663% | 0.7413%
| Z-Score (Absolute) |  | 0.39 | 0 | 0.63 | 0.29 | 1.88 | 0.72 | 0.9 | 0.89 | 2.01 | 0.67 | 0.14
| Analytical results for Tl in OREAS 462 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 14 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 17
| 1 | 3 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | < 0.5
| 1 | 5 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 23
| 1 | 6 | < 0.5 | < 0.5 | < 0.1 | < 0.5 | < 2 | < 0.5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Tm in OREAS 462 (Certified Value 1.2 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.13 | 1.11 | 1.16 | 1.24 | 1.21 | 1.1554 | 1.4 | 1.5 | 0.75649 | 1.2 | 1.2 | 0.94985 | 1.25 | 1.4 | 1.27 | 1
| 1 | 2 | 1.11 | 1.26 | 1.13 | 1.4 | 1.31 | 1.1578 | 1.2 | 2.5 | 0.812 | 1.12 | 1.3 | 0.96816 | 1.19 | 1.2 | 1.14 | 4
| 1 | 3 | 1.13 | 1.19 | 1.24 | 1.27 | 1.31 | 1.0886 | 1.2 | 1.4 | 0.78026 | 1.08 | 1.2 | 0.93786 | 1.17 | 1.4 | 0.9 | 16
| 1 | 4 | 1.18 | 1.24 | 1.17 | 1.25 | 1.28 | 1.1699 | 1.2 | 1.3 | 0.81858 | 1.09 | 1.1 | 0.9347 | 1.02 | 1.2 | 1.24 | 1.2039
| 1 | 5 | 1.21 | 1.2 | 1.19 | 1.21 | 1.23 | 1.0556 | 1.2 | 1.3 | 0.8034 | 1.03 | 1.2 | 0.91826 | 1.27 | 1.4 | 1.01 | 57
| 1 | 6 | 1.14 | 1.25 | 1.15 | 1.18 | 1.18 | 1.2647 | 1.2 | 1.3 | 0.85625 | 0.99 | 1.2 | 0.91468 | 1.22 | 1.4 | 0.97
| Mean |  | 1.15 | 1.2083 | 1.1733 | 1.2583 | 1.2533 | 1.1486 | 1.2333 | 1.55 | 0.8045 | 1.085 | 1.2 | 0.93725 | 1.1867 | 1.3333 | 1.0883
| Median |  | 1.135 | 1.22 | 1.165 | 1.245 | 1.255 | 1.1566 | 1.2 | 1.35 | 0.8077 | 1.085 | 1.2 | 0.93628 | 1.205 | 1.4 | 1.075
| Std Dev. |  | 0.037417 | 0.055648 | 0.038297 | 0.076267 | 0.05465 | 0.072538 | 0.08165 | 0.47223 | 0.034136 | 0.07287 | 0.063246 | 0.019955 | 0.089592 | 0.10328 | 0.15118
| Rel.Std.Dev. |  | 3.254% | 4.605% | 3.264% | 6.061% | 4.36% | 6.315% | 6.62% | 30.47% | 4.243% | 6.716% | 5.27% | 2.129% | 7.55% | 7.746% | 13.89%
| PDM3 |  | -4.475% | 0.3707% | -2.537% | 4.524% | 4.109% | -4.587% | 2.447% | 28.75% | -33.17% | -9.874% | -0.3215% | -22.15% | -1.429% | 10.75% | -9.597%
| Z-Score (Absolute) |  | 0.66 | 0.11 | 0.35 | 0.39 | 0.7 | 0.67 | 0 | 2.1 | 5.19 | 1.51 | 0 | 3.45 | 0.26 | 1.75 | 1.47
| Analytical results for U in OREAS 462 (Certified Value 6.15 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 5.91 | 5.79 | 6.52 | 6.08 | 6.52 | 6.2356 | 6.3 | 6.3 | 4.5109 | 6.31 | 6.5 | 5.5729 | 4.04 | 6.5 | 6.15 | 1
| 1 | 2 | 5.95 | 6.08 | 6.1 | 6.41 | 6.74 | 6.4785 | 5.9 | 6.3 | 4.8634 | 5.7 | 6.5 | 5.5444 | 3.96 | 6 | 6.56 | 5
| 1 | 3 | 5.65 | 6.1 | 6.37 | 6.35 | 6.04 | 6.0279 | 6.1 | 6.3 | 5.0364 | 5.68 | 6 | 5.5407 | 3.84 | 6 | 6.57 | 16
| 1 | 4 | 6.01 | 6.16 | 6.13 | 6 | 6.18 | 6.0807 | 5.9 | 6.3 | 4.9813 | 5.48 | 6.5 | 5.5105 | 3.3 | 6.5 | 6.62 | 6.1476
| 1 | 5 | 5.92 | 6.05 | 6.5 | 6.29 | 6.16 | 6.2996 | 6 | 6.3 | 5.171 | 5.52 | 6 | 5.3607 | 4.06 | 6 | 6.11 | 58
| 1 | 6 | 5.81 | 6.13 | 6.32 | 5.78 | 6.13 | 6.5009 | 6.1 | 6.3 | 5.1832 | 5.24 | 6.5 | 5.4232 | 3.92 | 6 | 5.82
| Mean |  | 5.875 | 6.0517 | 6.3233 | 6.1517 | 6.295 | 6.2706 | 6.05 | 6.3 | 4.9577 | 5.655 | 6.3333 | 5.4921 | 3.8533 | 6.1667 | 6.305
| Median |  | 5.915 | 6.09 | 6.345 | 6.185 | 6.17 | 6.2676 | 6.05 | 6.3 | 5.0088 | 5.6 | 6.5 | 5.5256 | 3.94 | 6 | 6.355
| Std Dev. |  | 0.12802 | 0.13378 | 0.17851 | 0.24145 | 0.27289 | 0.19661 | 0.15166 | 0 | 0.24964 | 0.36132 | 0.2582 | 0.082442 | 0.28275 | 0.2582 | 0.32611
| Rel.Std.Dev. |  | 2.179% | 2.211% | 2.823% | 3.925% | 4.335% | 3.135% | 2.507% | 0% | 5.035% | 6.389% | 4.077% | 1.501% | 7.338% | 4.187% | 5.172%
| PDM3 |  | -4.434% | -1.561% | 2.859% | 0.06612% | 2.398% | 2% | -1.588% | 2.479% | -19.36% | -8.013% | 3.021% | -10.66% | -37.32% | 0.3101% | 2.56%
| Z-Score (Absolute) |  | 1.22 | 0.44 | 0.75 | 0 | 0.24 | 0.52 | 0.45 | 0.65 | 5.25 | 1.82 | 0.8 | 2.9 | 9.62 | 0.07 | 0.67
| Analytical results for V in OREAS 462 (Certified Value 380 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 372 | 376 | 381 | 387 | 414 | 401 | 377 | 390 | 400 | 355.1091 | 367 | 400 | 360 |  |  | 1
| 1 | 2 | 369 | 387 | 390 | 405 | 430 | 395 | 362 | 357 | 400 | 359.2728 | 388 | 400 | 367 |  |  | 6
| 1 | 3 | 364 | 393 | 383 | 389 | 388 | 389 | 376 | 359 | 400 | 362.5247 | 371 | 400 | 368 |  |  | 16
| 1 | 4 | 364 | 389 | 388 | 389 | 377 | 386 | 387 | 352 | 400 | 368.7771 | 314 | 400 | 391 |  |  | 380.1715
| 1 | 5 | 372 | 379 | 394 | 398 | 382 | 385 | 370 | 369 | 400 | 357.8459 | 376 | 350 | 351 |  |  | 59
| 1 | 6 | 361 | 395 | 395 | 361 | 389 | 384 | 377 | 336 | 400 | 361.2458 | 388 | 350 | 338
| Mean |  | 367 | 386.5 | 388.5 | 388.1667 | 396.6667 | 390 | 374.8333 | 360.5 | 400 | 360.7959 | 367.3333 | 383.3333 | 362.5
| Median |  | 366.5 | 388 | 389 | 389 | 388.5 | 387.5 | 376.5 | 358 | 400 | 360.2593 | 373.5 | 400 | 363.5
| Std Dev. |  | 4.6476 | 7.5829 | 5.6833 | 14.9722 | 20.7236 | 6.6933 | 8.3287 | 18.0527 | 6.2269e-14 | 4.6958 | 27.5221 | 25.8199 | 17.8969
| Rel.Std.Dev. |  | 1.266% | 1.962% | 1.463% | 3.857% | 5.224% | 1.716% | 2.222% | 5.008% | 1.557e-14% | 1.302% | 7.492% | 6.736% | 4.937%
| PDM3 |  | -3.465% | 1.665% | 2.191% | 2.103% | 4.339% | 2.585% | -1.404% | -5.174% | 5.216% | -5.097% | -3.377% | 0.8317% | -4.648%
| Z-Score (Absolute) |  | 1.07 | 0.21 | 0.34 | 0.67 | 0.88 | 0.44 | 0.56 | 1.5 | 1.09 | 1.48 | 0.35 | 0 | 1.37
| Analytical results for W in OREAS 462 (Certified Value 3.6 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 17 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 5 | 3 | 4 | 4 | 4 | 4 | 3 | < 10 | 2 | < 5 | 8 |  |  |  |  | 1
| 1 | 2 | 6 | 4 | 4 | 4 | 5 | 4 | 3 | < 10 | 3 | < 5 | 9 |  |  |  |  | 7
| 1 | 3 | 4 | 4 | 4 | 3 | 5 | 3 | 3 | < 10 | 3 | < 5 | 8 |  |  |  |  | 16
| 1 | 4 | 4 | 4 | 4 | 4 | 4 | 3 | 3 | < 10 | 3 | < 5 | 8 |  |  |  |  | 3.5952
| 1 | 5 | 7 | 4 | 4 | 4 | 5 | 3 | 3 | < 10 | 3 | < 5 | 8 |  |  |  |  | 60
| 1 | 6 | 6 | 4 | 4 | 3 | 4 | 3 | 3 | < 10 | 3 | < 5 | 8
| Mean |  | 5.3333 | 3.8333 | 4 | 3.6667 | 4.5 | 3.3333 | 3 |  | 2.8333 |  | 8.1667
| Median |  | 5.5 | 4 | 4 | 4 | 4.5 | 3 | 3 |  | 3 |  | 8
| Std Dev. |  | 1.2111 | 0.40825 | 9.7295e-16 | 0.5164 | 0.54772 | 0.5164 | 4.8648e-16 |  | 0.40825 |  | 0.40825
| Rel.Std.Dev. |  | 22.71% | 10.65% | 2.432e-14% | 14.08% | 12.17% | 15.49% | 1.622e-14% |  | 14.41% |  | 4.999%
| PDM3 |  | 48.34% | 6.623% | 11.26% | 1.987% | 25.17% | -7.285% | -16.56% |  | -21.19% |  | 127.2%
| Z-Score (Absolute) |  | 1.52 | 0 | 0.17 | 0.17 | 0.67 | 0.51 | 0.84 | NA | 1.01 | NA | 4.38
| Analytical results for Y in OREAS 462 (Certified Value 133 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 129.5 | 140 | 140 | 134.5 | 135 | 132.0381 | 138.8 | 132 | 128.68 | 130.5 | 132.7687 | 202 | 134 | 123 |  | 1
| 1 | 2 | 130.5 | 140 | 141.5 | 141 | 139.5 | 131.9974 | 132.9 | 127.9 | 120.39 | 129 | 134.0302 | 203 | 135 | 129 |  | 8
| 1 | 3 | 126.5 | 142 | 141.5 | 137.5 | 141 | 131.9237 | 131.1 | 132 | 118.29 | 127 | 135.6478 | 205 | 134 | 128 |  | 16
| 1 | 4 | 128 | 139.5 | 141 | 135.5 | 139.5 | 134.0635 | 132.6 | 132.9 | 118.85 | 129 | 131.7596 | 205 | 135 | 137 |  | 132.6531
| 1 | 5 | 130.5 | 137 | 143.5 | 139 | 142.5 | 131.6648 | 135.7 | 132.6 | 115.91 | 128 | 128.6811 | 195 | 135 | 120 |  | 61
| 1 | 6 | 125.5 | 143 | 143.5 | 127 | 140 | 130.756 | 132.7 | 130.2 | 109.97 | 128.5 | 128.6655 | 200 | 134 | 118
| Mean |  | 128.4167 | 140.25 | 141.8333 | 135.75 | 139.5833 | 132.0739 | 133.9667 | 131.2667 | 118.6817 | 128.6667 | 131.9255 | 201.6667 | 134.5 | 125.8333
| Median |  | 128.75 | 140 | 141.5 | 136.5 | 139.75 | 131.9605 | 132.8 | 132 | 118.57 | 128.75 | 132.2641 | 202.5 | 134.5 | 125.5
| Std Dev. |  | 2.1075 | 2.0917 | 1.4024 | 4.8862 | 2.5183 | 1.0856 | 2.7998 | 1.897 | 6.1097 | 1.169 | 2.8347 | 3.7771 | 0.54772 | 6.969
| Rel.Std.Dev. |  | 1.641% | 1.491% | 0.9888% | 3.599% | 1.804% | 0.822% | 2.09% | 1.445% | 5.148% | 0.9086% | 2.149% | 1.873% | 0.4072% | 5.538%
| PDM3 |  | -3.194% | 5.727% | 6.92% | 2.335% | 5.224% | -0.4366% | 0.9902% | -1.045% | -10.53% | -3.005% | -0.5485% | 52.03% | 1.392% | -5.141%
| Z-Score (Absolute) |  | 0.69 | 1.09 | 1.33 | 0.41 | 0.99 | 0.14 | 0.14 | 0.26 | 1.9 | 0.66 | 0.16 | 10.34 | 0.22 | 1.08
| Analytical results for Yb in OREAS 462 (Certified Value 5.51 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 5.46 | 5.79 | 5.7 | 5.87 | 5.61 | 5.0986 | 6 | 5.9 | 4.9998 | 5.3 | 5.8 | 5.4609 | 5.3 | 6 | 6 | 1
| 1 | 2 | 5.41 | 5.56 | 5.3 | 5.56 | 5.5 | 4.8564 | 5.2 | 5.9 | 5.2748 | 5 | 5.9 | 5.4137 | 6.6 | 5.5 | 6.7 | 9
| 1 | 3 | 5.32 | 5.46 | 5.9 | 5.35 | 5.44 | 5.4811 | 5.6 | 5.8 | 5.2184 | 4.9 | 6.4 | 5.4212 | 6.7 | 6.5 | 7.3 | 16
| 1 | 4 | 5.51 | 5.54 | 6.5 | 5.41 | 5.62 | 5.2698 | 5.7 | 5.6 | 5.2409 | 5.1 | 5.8 | 5.4438 | 6.7 | 5 | 7.1 | 5.5122
| 1 | 5 | 5.5 | 5.69 | 6.1 | 5.71 | 5.82 | 4.8372 | 5.7 | 5.7 | 5.2899 | 4.8 | 5.8 | 5.2808 | 6.4 | 5 | 6.1 | 62
| 1 | 6 | 5.3 | 5.6 | 5.9 | 5.28 | 5.63 | 5.2924 | 5.6 | 6 | 5.1415 | 4.7 | 5.8 | 5.1669 | 6.4 | 6 | 6
| Mean |  | 5.4167 | 5.6067 | 5.9 | 5.53 | 5.6033 | 5.1392 | 5.6333 | 5.8167 | 5.1942 | 4.9667 | 5.9167 | 5.3645 | 6.35 | 5.6667 | 6.5333
| Median |  | 5.435 | 5.58 | 5.9 | 5.485 | 5.615 | 5.1842 | 5.65 | 5.85 | 5.2297 | 4.95 | 5.8 | 5.4175 | 6.5 | 5.75 | 6.4
| Std Dev. |  | 0.090037 | 0.11725 | 0.4 | 0.22724 | 0.13064 | 0.25698 | 0.2582 | 0.1472 | 0.10861 | 0.21602 | 0.24014 | 0.11598 | 0.53198 | 0.60553 | 0.58195
| Rel.Std.Dev. |  | 1.662% | 2.091% | 6.78% | 4.109% | 2.331% | 5% | 4.583% | 2.531% | 2.091% | 4.349% | 4.059% | 2.162% | 8.378% | 10.69% | 8.907%
| PDM3 |  | -1.732% | 1.715% | 7.036% | 0.3237% | 1.654% | -6.765% | 2.198% | 5.524% | -5.768% | -9.896% | 7.339% | -2.678% | 15.2% | 2.803% | 18.53%
| Z-Score (Absolute) |  | 0.6 | 0 | 0.93 | 0.24 | 0.01 | 1.48 | 0.08 | 0.66 | 1.3 | 2.02 | 0.67 | 0.77 | 3.01 | 0.19 | 2.93
| Analytical results for Zn in OREAS 462 (Certified Value 269 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 14 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 255 | 300 | 369.9173 | 300 | 250 | 241 | [9 empty] | 1
| 1 | 2 | 259 | 300 | 358.6288 | 300 | 250 | 246 | [9 empty] | 0
| 1 | 3 | 250 | 300 | 366.9556 | 320 | 250 | 243 | [9 empty] | 16
| 1 | 4 | 258 | 300 | 351.3236 | 300 | 250 | 243 | [9 empty] | 268.9
| 1 | 5 | 242 | 300 | 336.2672 | 300 | 250 | 242 | [9 empty] | 63
| 1 | 6 | 241 | 300 | 366.44 | 300 | 250 | 247
| Mean |  | 250.8333 | 300 | 358.2554 | 303.3333 | 250 | 243.6667
| Median |  | 252.5 | 300 | 362.5344 | 300 | 250 | 243
| Std Dev. |  | 7.8846 | 6.2269e-14 | 12.7307 | 8.165 | 0 | 2.3381
| Rel.Std.Dev. |  | 3.143% | 2.076e-14% | 3.554% | 2.692% | 0% | 0.9595%
| PDM3 |  | -6.719% | 11.57% | 33.23% | 12.81% | -7.029% | -9.384%
| Z-Score (Absolute) |  | 0.66 | 0.66 | 2.23 | 0.66 | 0.69 | 0.86
| Analytical results for Zr in OREAS 462 (Certified Value 593 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 581 | 613 | 622 | 614 | 605 | 600 | 557 | 595 | 561 | 671.5 | 592 |  |  |  |  | 1
| 1 | 2 | 587 | 619 | 617 | 645 | 632 | 585 | 575 | 604 | 567 | 580.4 | 581 |  |  |  |  | 10
| 1 | 3 | 568 | 637 | 611 | 635 | 568 | 573 | 586 | 597 | 572 | 546.3 | 645 |  |  |  |  | 16
| 1 | 4 | 581 | 619 | 619 | 637 | 584 | 582 | 569 | 600 | 564 | 561.9 | 610 |  |  |  |  | 593.3136
| 1 | 5 | 584 | 613 | 626 | 618 | 599 | 589 | 581 | 584 | 583 | 530.1 | 619 |  |  |  |  | 64
| 1 | 6 | 579 | 641 | 638 | 589 | 567 | 580 | 577 | 578 | 557 | 548.6 | 627
| Mean |  | 580 | 623.6667 | 622.1667 | 623 | 592.5 | 584.8333 | 574.1667 | 593 | 567.3333 | 573.1333 | 612.3333
| Median |  | 581 | 619 | 620.5 | 626.5 | 591.5 | 583.5 | 576 | 596 | 565.5 | 555.25 | 614.5
| Std Dev. |  | 6.5115 | 12.242 | 9.2394 | 20.4255 | 24.8254 | 9.1524 | 10.1669 | 9.9599 | 9.2232 | 51.0466 | 23.3552
| Rel.Std.Dev. |  | 1.123% | 1.963% | 1.485% | 3.279% | 4.19% | 1.565% | 1.771% | 1.68% | 1.626% | 8.907% | 3.814%
| PDM3 |  | -2.244% | 5.116% | 4.863% | 5.003% | -0.1371% | -1.429% | -3.227% | -0.05286% | -4.379% | -3.401% | 3.206%
| Z-Score (Absolute) |  | 0.42 | 1.06 | 1.01 | 1.04 | 0 | 0.26 | 0.62 | 0.02 | 0.86 | 1.33 | 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 462 (Certified Value < 2 ppm) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
2
| 1 | 1 | 1.07 | 1.57 | 2.85 | 0.57 | 1.07 | 0.89554 | 1.2 | 1.4 | < 0.01 | 1.88 | < 1 | 1.5 |  |  | 1
| 1 | 2 | 0.73 | 1.49 | 0.72 | 0.69 | 1 | 0.96629 | 1.1 | 1.1 | < 0.01 | 1.47 | < 1 | 1.5 |  |  | 15
| 1 | 3 | 1.16 | 1.78 | 1.76 | 0.59 | 0.84 | 0.93002 | 1 | 1.2 | < 0.01 | 1.69 | 1 | 1.5 |  |  | 16
| 1 | 4 | 1.11 | 1.53 | 1.33 | 0.61 | 0.74 | 0.86341 | 1.4 | 1.2 | < 0.01 | 1.69 | 2 | 1.5 |  |  | < 2
| 1 | 5 | 1.86 | 1.38 | 1.42 | 0.6 | 0.71 | 0.96602 | 1.2 | 1.3 | < 0.01 | 1.91 | 1 | 1.5 |  |  | 66
| 1 | 6 | 1.11 | 1.5 | 0.05 | 0.61 | 0.73 | 0.97 | 1.2 | 1.2 | < 0.01 | 1.66 | 1 | 1.5
| Mean |  | 1.1733 | 1.5417 | 1.355 | 0.61167 | 0.84833 | 0.93188 | 1.1833 | 1.2333 |  | 1.7167 | 1.25 | 1.5
| Median |  | 1.11 | 1.515 | 1.375 | 0.605 | 0.79 | 0.94802 | 1.2 | 1.2 |  | 1.69 | 1 | 1.5
| Std Dev. |  | 0.37066 | 0.13288 | 0.94975 | 0.041191 | 0.15303 | 0.044306 | 0.13292 | 0.10328 |  | 0.16096 | 0.5 | 0
| Rel.Std.Dev. |  | 31.59% | 8.619% | 70.09% | 6.734% | 18.04% | 4.754% | 11.23% | 8.374% |  | 9.376% | 40% | 0%
PDM3
| Z-Score (Absolute) |  | 0.4 | 0.96 | 0.34 | 1.53 | 0.9 | 0.67 | 0 | 0.13 | NA | 1.43 | 0.18 | 0.85
| Analytical results for Al in OREAS 462 (Certified Value 5.71 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 5.28 | 5.86 | 5.5 | 5.9 | 5.32 | 5.7061 | 5.72 | 5.79 | 5.75 | 5.6421 | 5.6849 | 5.79 | 5.912 | 6.09 | 1
| 1 | 2 | 5.41 | 5.86 | 5.51 | 5.93 | 5.57 | 5.739 | 5.7 | 5.73 | 5.73 | 5.8113 | 5.6185 | 5.73 | 5.862 | 5.99 | 1
| 1 | 3 | 4.59 | 6.04 | 5.48 | 5.9 | 5.7 | 5.5857 | 5.67 | 5.76 | 5.7 | 5.5767 | 5.6837 | 5.69 | 5.832 | 6.01 | 16
| 1 | 4 | 5 | 5.81 | 5.58 | 5.95 | 5.28 | 5.7249 | 5.75 | 5.7 | 5.77 | 5.6762 | 5.7422 | 5.74 | 5.884 | 6.02 | 5.7069
| 1 | 5 | 5.11 | 5.78 | 5.48 | 5.72 | 6.5 | 5.6262 | 5.72 | 5.65 | 5.67 | 5.5736 | 5.6296 | 5.76 | 5.859 | 6.17 | 67
| 1 | 6 | 4.59 | 5.85 | 5.59 | 5.99 | 5.43 | 5.5345 | 5.69 | 5.66 | 5.68 | 5.5461 | 5.6181 | 5.65 | 5.862 | 6.06
| Mean |  | 4.9967 | 5.8667 | 5.5233 | 5.8983 | 5.6333 | 5.6527 | 5.7083 | 5.715 | 5.7167 | 5.6377 | 5.6628 | 5.7267 | 5.8685 | 6.0567
| Median |  | 5.055 | 5.855 | 5.505 | 5.915 | 5.5 | 5.6661 | 5.71 | 5.715 | 5.715 | 5.6094 | 5.6566 | 5.735 | 5.862 | 6.04
| Std Dev. |  | 0.34489 | 0.090701 | 0.049261 | 0.093684 | 0.45253 | 0.083267 | 0.027869 | 0.055408 | 0.039833 | 0.097836 | 0.049586 | 0.050067 | 0.02698 | 0.066232
| Rel.Std.Dev. |  | 6.902% | 1.546% | 0.8919% | 1.588% | 8.033% | 1.473% | 0.4882% | 0.9695% | 0.6968% | 1.735% | 0.8756% | 0.8743% | 0.4597% | 1.094%
| PDM3 |  | -12.45% | 2.8% | -3.216% | 3.355% | -1.289% | -0.9493% | 0.02526% | 0.1421% | 0.1713% | -1.213% | -0.7723% | 0.3465% | 2.832% | 6.129%
| Z-Score (Absolute) |  | 4.96 | 0.84 | 1.31 | 1.54 | 1.43 | 0.41 | 0.02 | 0.02 | 0.03 | 0.51 | 0.34 | 0.1 | 1.09 | 2.39
| Analytical results for As in OREAS 462 (Certified Value 30.8 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 30.8 | 31.8 | 31.9 | 30.8 | 31.3 | 28.5 | 31 | 31 | 30.9 | 52.9 | < 50 | 38 |  |  | 1
| 1 | 2 | 31.9 | 31.2 | 31 | 30.1 | 32 | 27.7952 | 31 | 32 | 30.4 | 55.4 | < 50 | 39 |  |  | 16
| 1 | 3 | 32.3 | 34.5 | 32.9 | 30 | 31 | 27.6241 | 31 | 33 | 30 | 56.5 | < 50 | 39 |  |  | 16
| 1 | 4 | 31 | 31.9 | 31.3 | 29.5 | 30.3 | 28.9364 | 31 | 33 | 30.2 | 48.3 | < 50 | 38 |  |  | 30.7827
| 1 | 5 | 30.6 | 31 | 33.1 | 30.2 | 30 | 27.9487 | 30 | 30 | 32 | 47 | < 50 | 38 |  |  | 68
| 1 | 6 | 31 | 31.2 | 31.1 | 30.4 | 30.7 | 27.8397 | 31 | 36 | 30.6 | 51.6 | 50 | 39
| Mean |  | 31.2667 | 31.9333 | 31.8833 | 30.1667 | 30.8833 | 28.1073 | 30.8333 | 32.5 | 30.6833 | 51.95 | 50 | 38.5
| Median |  | 31 | 31.5 | 31.6 | 30.15 | 30.85 | 27.8942 | 31 | 32.5 | 30.5 | 52.25 | 50 | 38.5
| Std Dev. |  | 0.67429 | 1.3079 | 0.92177 | 0.43205 | 0.71949 | 0.50385 | 0.40825 | 2.0736 | 0.71671 | 3.7803 |  | 0.54772
| Rel.Std.Dev. |  | 2.157% | 4.096% | 2.891% | 1.432% | 2.33% | 1.793% | 1.324% | 6.38% | 2.336% | 7.277% |  | 1.423%
| PDM3 |  | 1.572% | 3.738% | 3.576% | -2.001% | 0.327% | -8.691% | 0.1646% | 5.579% | -0.3227% | 68.76% | 62.43% | 25.07%
| Z-Score (Absolute) |  | 0.09 | 0.09 | 0.61 | 1.32 | 0.52 | 3.64 | 0.57 | 0.51 | 0.74 | 23.16 | 20.97 | 8.04
| Analytical results for Ba in OREAS 462 (Certified Value 1037 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 990 | 1050 | 1060 | 1090 | 1000 | 1034.4228 | 1029 | 1020 | 1130 | 1016.7 | 1083.37 | 390 | 1035 | 974 | 1
| 1 | 2 | 990 | 1060 | 1060 | 1110 | 1030 | 1025.2731 | 1024 | 1019 | 1110 | 1050.6 | 1106.63 | 391 | 1030 | 1020 | 2
| 1 | 3 | 950 | 1070 | 1040 | 1100 | 1080 | 1023.4178 | 1004 | 1031 | 1120 | 1007 | 1102.37 | 394 | 1035 | 1020 | 16
| 1 | 4 | 970 | 1040 | 1070 | 1100 | 1000 | 1031.9371 | 1024 | 1030 | 1140 | 1066 | 1116.17 | 392 | 1030 | 1010 | 1036.7978
| 1 | 5 | 1010 | 1040 | 1050 | 1080 | 1210 | 1018.612 | 1017 | 997 | 1130 | 1009.6 | 1086.06 | 394 | 1030 | 1020 | 69
| 1 | 6 | 930 | 1050 | 1070 | 1100 | 1010 | 1028.5664 | 996 | 1006 | 1140 | 1055.8 | 1089.91 | 385 | 1020 | 1010
| Mean |  | 973.3333 | 1051.6667 | 1058.3333 | 1096.6667 | 1055 | 1027.0382 | 1015.6667 | 1017.1667 | 1128.3333 | 1034.2833 | 1097.4183 | 391 | 1030 | 1009
| Median |  | 980 | 1050 | 1060 | 1100 | 1020 | 1026.9197 | 1020.5 | 1019.5 | 1130 | 1033.65 | 1096.14 | 391.5 | 1030 | 1015
| Std Dev. |  | 29.4392 | 11.6905 | 11.6905 | 10.328 | 81.6701 | 5.7981 | 12.9718 | 13.4077 | 11.6905 | 26.069 | 12.9903 | 3.3466 | 5.4772 | 17.8326
| Rel.Std.Dev. |  | 3.025% | 1.112% | 1.105% | 0.9418% | 7.741% | 0.5645% | 1.277% | 1.318% | 1.036% | 2.52% | 1.184% | 0.8559% | 0.5318% | 1.767%
| PDM3 |  | -6.121% | 1.434% | 2.077% | 5.774% | 1.756% | -0.9413% | -2.038% | -1.893% | 8.829% | -0.2425% | 5.847% | -62.29% | -0.6557% | -2.681%
| Z-Score (Absolute) |  | 2.07 | 0.87 | 1.12 | 2.55 | 0.17 | 0.06 | 0.48 | 0.43 | 3.74 | 0.22 | 2.58 | 23.88 | 0.06 | 0.47
| Analytical results for Be in OREAS 462 (Certified Value 3.57 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.94 | 3.51 | 3.18 | 4.03 | 3.02 | 3.6997 | 3.8 | 4.2 | 3.42 | 3.77 | 3.1455 | 4.5 | 3.6 |  | 1
| 1 | 2 | 2.81 | 3.76 | 2.9 | 3.95 | 3.26 | 3.6873 | 3.8 | 3.5 | 3.38 | 3.88 | 3.1624 | 4.5 | 3.9 |  | 0
| 1 | 3 | 2.86 | 4.08 | 3.31 | 3.95 | 3.08 | 3.5518 | 3.7 | 3.5 | 3.41 | 4.2 | 3.1168 | 4.5 | 3.7 |  | 16
| 1 | 4 | 2.75 | 3.71 | 3.19 | 3.72 | 3 | 3.5459 | 3.6 | 4 | 3.43 | 3.55 | 3.156 | 5 | 3.8 |  | 3.5673
| 1 | 5 | 2.84 | 3.81 | 3.56 | 3.73 | 3.08 | 3.7082 | 3.8 | 4 | 3.34 | 3.86 | 3.1153 | 4.5 | 3.7 |  | 70
| 1 | 6 | 2.81 | 3.95 | 3.17 | 3.87 | 2.96 | 3.6453 | 3.7 | 3.7 | 3.38 | 3.61 | 3.1999 | 4.5 | 3.7
| Mean |  | 2.835 | 3.8033 | 3.2183 | 3.875 | 3.0667 | 3.6397 | 3.7333 | 3.8167 | 3.3933 | 3.8117 | 3.1493 | 4.5833 | 3.7333
| Median |  | 2.825 | 3.785 | 3.185 | 3.91 | 3.05 | 3.6663 | 3.75 | 3.85 | 3.395 | 3.815 | 3.1508 | 4.5 | 3.7
| Std Dev. |  | 0.063482 | 0.19735 | 0.21498 | 0.12677 | 0.10558 | 0.073633 | 0.08165 | 0.29269 | 0.033267 | 0.23181 | 0.031618 | 0.20412 | 0.10328
| Rel.Std.Dev. |  | 2.239% | 5.189% | 6.68% | 3.271% | 3.443% | 2.023% | 2.187% | 7.669% | 0.9804% | 6.082% | 1.004% | 4.454% | 2.766%
| PDM3 |  | -20.53% | 6.616% | -9.783% | 8.624% | -14.03% | 2.029% | 4.653% | 6.989% | -4.878% | 6.849% | -11.72% | 28.48% | 4.653%
| Z-Score (Absolute) |  | 4.28 | 0.33 | 2.45 | 0.67 | 3.17 | 0.45 | 0 | 0.4 | 1.62 | 0.37 | 2.78 | 3.65 | 0
| Analytical results for Bi in OREAS 462 (Certified Value 2.69 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 2.87 | 2.72 | 3.01 | 2.74 | 2.42 | 2.6798 | 2.9 | 2.9 | < 0.01 | 2.76 | 2.3757 | 2.4 | 2.5 |  | 1
| 1 | 2 | 2.85 | 2.76 | 2.78 | 2.69 | 2.6 | 2.6668 | 2.8 | 2.9 | < 0.01 | 2.84 | 2.4903 | 2.7 | 2.8 |  | 0
| 1 | 3 | 2.86 | 2.94 | 3 | 2.76 | 2.47 | 2.5691 | 2.8 | 2.9 | < 0.01 | 2.76 | 2.4897 | 2.5 | 2.6 |  | 16
| 1 | 4 | 2.85 | 2.7 | 2.6 | 2.71 | 2.37 | 2.4939 | 2.9 | 2.9 | < 0.01 | 2.69 | 2.4401 | 2.5 | 2.5 |  | 2.6935
| 1 | 5 | 2.79 | 2.74 | 2.96 | 2.66 | 2.38 | 2.6246 | 2.8 | 3 | < 0.01 | 2.74 | 2.5405 | 2.4 | 2.6 |  | 71
| 1 | 6 | 2.73 | 2.83 | 2.82 | 2.79 | 2.47 | 2.6365 | 2.8 | 2.9 | < 0.01 | 2.73 | 2.4253 | 2.7 | 2.6
| Mean |  | 2.825 | 2.7817 | 2.8617 | 2.725 | 2.4517 | 2.6118 | 2.8333 | 2.9167 |  | 2.7533 | 2.4603 | 2.5333 | 2.6
| Median |  | 2.85 | 2.75 | 2.89 | 2.725 | 2.445 | 2.6305 | 2.8 | 2.9 |  | 2.75 | 2.4649 | 2.5 | 2.6
| Std Dev. |  | 0.054314 | 0.089536 | 0.1598 | 0.047645 | 0.084242 | 0.069481 | 0.05164 | 0.040825 |  | 0.049666 | 0.058278 | 0.13663 | 0.10954
| Rel.Std.Dev. |  | 1.923% | 3.219% | 5.584% | 1.748% | 3.436% | 2.66% | 1.823% | 1.4% |  | 1.804% | 2.369% | 5.393% | 4.213%
| PDM3 |  | 4.882% | 3.273% | 6.243% | 1.169% | -8.979% | -3.034% | 5.191% | 8.285% |  | 2.221% | -8.659% | -5.947% | -3.472%
| Z-Score (Absolute) |  | 0.47 | 0.07 | 0.67 | 0.07 | 1.54 | 0.68 | 0.52 | 0.97 | NA | 0.09 | 1.5 | 1.1 | 0.74
| Analytical results for Ca in OREAS 462 (Certified Value 1.25 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.24 | 1.28 | 0.97 | 1.01 | 0.92 | 1.305 | 1.23 | 1.36 | 1.16 | 1.2509 | 1.2492 | 1.23 | 1.2775 | 1.38 | 1
| 1 | 2 | 1.24 | 1.26 | 0.97 | 1.05 | 0.93 | 1.2923 | 1.22 | 1.33 | 1.17 | 1.3673 | 1.2464 | 1.22 | 1.279 | 1.37 | 3
| 1 | 3 | 1.18 | 1.31 | 0.97 | 1.07 | 0.93 | 1.301 | 1.22 | 1.33 | 1.16 | 1.3256 | 1.2527 | 1.21 | 1.2785 | 1.38 | 16
| 1 | 4 | 1.2 | 1.26 | 0.98 | 1.06 | 0.94 | 1.3084 | 1.24 | 1.31 | 1.17 | 1.3365 | 1.2604 | 1.22 | 1.268 | 1.4 | 1.2475
| 1 | 5 | 1.21 | 1.24 | 0.97 | 1.03 | 1.02 | 1.3125 | 1.23 | 1.3 | 1.15 | 1.2728 | 1.2558 | 1.22 | 1.271 | 1.45 | 72
| 1 | 6 | 1.17 | 1.25 | 0.98 | 1.02 | 0.96 | 1.2814 | 1.23 | 1.3 | 1.15 | 1.3182 | 1.2458 | 1.2 | 1.3105 | 1.4
| Mean |  | 1.2067 | 1.2667 | 0.97333 | 1.04 | 0.95 | 1.3001 | 1.2283 | 1.3217 | 1.16 | 1.3119 | 1.2517 | 1.2167 | 1.2808 | 1.3967
| Median |  | 1.205 | 1.26 | 0.97 | 1.04 | 0.935 | 1.303 | 1.23 | 1.32 | 1.16 | 1.3219 | 1.2509 | 1.22 | 1.278 | 1.39
| Std Dev. |  | 0.029439 | 0.025033 | 0.005164 | 0.023664 | 0.036878 | 0.011477 | 0.0075277 | 0.023166 | 0.0089443 | 0.042782 | 0.0056895 | 0.010328 | 0.015241 | 0.028752
| Rel.Std.Dev. |  | 2.44% | 1.976% | 0.5305% | 2.275% | 3.882% | 0.8828% | 0.6128% | 1.753% | 0.7711% | 3.261% | 0.4545% | 0.8489% | 1.19% | 2.059%
| PDM3 |  | -3.276% | 1.533% | -21.98% | -16.64% | -23.85% | 4.213% | -1.539% | 5.942% | -7.017% | 5.158% | 0.3351% | -2.475% | 2.662% | 11.95%
| Z-Score (Absolute) |  | 0.34 | 0.27 | 2.73 | 2.04 | 3.11 | 0.61 | 0.12 | 0.83 | 0.82 | 0.73 | 0.12 | 0.24 | 0.42 | 1.49
| Analytical results for Cd in OREAS 462 (Certified Value < 0.2 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.08 | 0.05 | 0.03 | 0.11 | 0.11 | 0.082205 | < 0.2 | 0.47 | 0.04 | < 0.5 | 0.5 |  |  |  | 1
| 1 | 2 | 0.11 | 0.11 | 0.12 | 0.12 | 0.11 | 0.05 | < 0.2 | 0.47 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 18
| 1 | 3 | 0.07 | 0.09 | 0.15 | 0.03 | 0.09 | 0.08 | 0.3 | 0.5 | 0.08 | < 0.5 | 0.5 |  |  |  | 16
| 1 | 4 | 0.19 | 0.12 | 0.15 | 0.1 | 0.1 | 0.08 | < 0.2 | 0.49 | < 0.02 | < 0.5 | < 0.5 |  |  |  | < 0.2
| 1 | 5 | 0.16 | 0.12 | 0.17 | < 0.02 | 0.12 | 0.047738 | 0.2 | 0.43 | < 0.02 | < 0.5 | < 0.5 |  |  |  | 73
| 1 | 6 | 0.08 | 0.1 | 0.04 | 0.13 | 0.1 | 0.09 | < 0.2 | 0.49 | < 0.02 | < 0.5 | 0.5
| Mean |  | 0.115 | 0.098333 | 0.11 | 0.098 | 0.105 | 0.071657 | 0.25 | 0.475 | 0.06 |  | 0.5
| Median |  | 0.095 | 0.105 | 0.135 | 0.11 | 0.105 | 0.08 | 0.25 | 0.48 | 0.06 |  | 0.5
| Std Dev. |  | 0.049295 | 0.026394 | 0.060332 | 0.039623 | 0.010488 | 0.018045 | 0.070711 | 0.0251 | 0.028284 |  | 0
| Rel.Std.Dev. |  | 42.87% | 26.84% | 54.85% | 40.43% | 9.989% | 25.18% | 28.28% | 5.284% | 47.14% |  | 0%
PDM3
| Z-Score (Absolute) |  | 0.06 | 0.35 | 0.06 | 0.06 | 0.18 | 1 | 3.37 | 8.89 | 1.29 | NA | 9.5
| Analytical results for Ce in OREAS 462 (Certified Value 4984 ppm) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 5132.9946 | 4955.8 | 4775.7 | 4980 | 4569.52 | 5227.6465 | 5222 | 4930 |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 5172.24 | 4873.7 | 4791.5 | 4940 | 4753.92 | 5182.134 | 5205 | 5050 |  | 0
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 5121.82 | 4817.1 | 4726.9 | 4930 | 4717.52 | 5163.469 | 5211 | 5100 |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 5071.7095 | 4823.7 | 4847.6 | 4940 | 4731.5 | 5130.4331 | 5251.5 | 4970 |  | 4984.2902
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 5119.8276 | 4901.8 | 4807.1 | 4910 | 4632.52 | 5235.5671 | 5192 | 5070 |  | 74
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 5173.92 | 4840.2 | 4793 | 4930 | 4722.66 | 5090.3234 | 5308.5 | 5060
| Mean | [6 empty] | 5132.0853 | 4868.7167 | 4790.3 | 4938.3333 | 4687.94 | 5171.5955 | 5231.6667 | 5030
| Median | [6 empty] | 5127.4073 | 4856.95 | 4792.25 | 4935 | 4720.09 | 5172.8015 | 5216.5 | 5055
| Std Dev. | [6 empty] | 38.1162 | 53.3045 | 39.5145 | 23.1661 | 71.3105 | 56.0547 | 42.6658 | 65.4217
| Rel.Std.Dev. | [6 empty] | 0.7427% | 1.095% | 0.8249% | 0.4691% | 1.521% | 1.084% | 0.8155% | 1.301%
| PDM3 | [6 empty] | 2.965% | -2.319% | -3.892% | -0.922% | -5.946% | 3.758% | 4.963% | 0.9171%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.59 | 0.46 | 0.78 | 0.18 | 1.1 | 0.75 | 1 | 0.18
| Analytical results for Co in OREAS 462 (Certified Value 13 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 12.7 | 12.6 | 13.3 | 13 | 12.3 | 11.6475 | 13.5 | 14 | 12.9 | 14 | 13.3699 | 25 | 15 |  | 1
| 1 | 2 | 12.6 | 12.6 | 12.3 | 13.1 | 13.1 | 11.7584 | 13.7 | 13 | 12.5 | 13.8 | 13.1093 | 25 | 15 |  | 0
| 1 | 3 | 12.7 | 14 | 13.7 | 13.1 | 12.6 | 11.8884 | 13.7 | 14 | 12.4 | 13.7 | 13.1983 | 25 | 15 |  | 16
| 1 | 4 | 12.4 | 12.6 | 13.1 | 12.8 | 12.3 | 11.264 | 13.8 | 13 | 12.5 | 13 | 13.1001 | 25 | 15 |  | 13.0237
| 1 | 5 | 12.1 | 12.9 | 13.3 | 12.7 | 12.2 | 11.5 | 13.6 | 13 | 12.4 | 13.4 | 13.1541 | 25 | 15 |  | 75
| 1 | 6 | 12.1 | 12.9 | 12.5 | 13.2 | 12.4 | 11.2994 | 13.7 | 13 | 12.6 | 13.2 | 13.2559 | 25 | 15
| Mean |  | 12.4333 | 12.9333 | 13.0333 | 12.9833 | 12.4833 | 11.5596 | 13.6667 | 13.3333 | 12.55 | 13.5167 | 13.1979 | 25 | 15
| Median |  | 12.5 | 12.75 | 13.2 | 13.05 | 12.35 | 11.5737 | 13.7 | 13 | 12.5 | 13.55 | 13.1762 | 25 | 15
| Std Dev. |  | 0.28048 | 0.54283 | 0.53166 | 0.19408 | 0.33116 | 0.25061 | 0.10328 | 0.5164 | 0.18708 | 0.38166 | 0.10228 | 0 | 0
| Rel.Std.Dev. |  | 2.256% | 4.197% | 4.079% | 1.495% | 2.653% | 2.168% | 0.7557% | 3.873% | 1.491% | 2.824% | 0.775% | 0% | 0%
| PDM3 |  | -4.533% | -0.6938% | 0.07408% | -0.3098% | -4.149% | -11.24% | 4.937% | 2.378% | -3.637% | 3.785% | 1.338% | 91.96% | 15.17%
| Z-Score (Absolute) |  | 0.73 | 0.38 | 0 | 0.06 | 0.82 | 1.8 | 0.77 | 0.37 | 0.67 | 0.59 | 0.2 | 14.58 | 2.4
| Analytical results for Cr in OREAS 462 (Certified Value 502 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 495 | 555 | 394 | 441 | 407 | 508.8447 | 556 | 502 | 651 | 391 | 600 |  |  |  | 1
| 1 | 2 | 495 | 547 | 395 | 459 | 406 | 512.3426 | 545 | 507 | 692 | 391 | 600 |  |  |  | 4
| 1 | 3 | 484 | 563 | 396 | 466 | 410 | 520.5729 | 540 | 512 | 652 | 400 | 650 |  |  |  | 16
| 1 | 4 | 481 | 543 | 401 | 459 | 412 | 536.6505 | 551 | 524 | 631 | 386 | 600 |  |  |  | 501.5086
| 1 | 5 | 493 | 537 | 394 | 450 | 443 | 527.326 | 553 | 512 | 622 | 393 | 600 |  |  |  | 76
| 1 | 6 | 477 | 540 | 394 | 451 | 416 | 507.6298 | 544 | 508 | 622 | 380 | 600
| Mean |  | 487.5 | 547.5 | 395.6667 | 454.3333 | 415.6667 | 518.8944 | 548.1667 | 510.8333 | 645 | 390.1667 | 608.3333
| Median |  | 488.5 | 545 | 394.5 | 455 | 411 | 516.4578 | 548 | 510 | 641 | 391 | 600
| Std Dev. |  | 7.8422 | 9.8336 | 2.7325 | 8.8015 | 13.866 | 11.4892 | 6.1128 | 7.4409 | 26.6383 | 6.7355 | 20.4124
| Rel.Std.Dev. |  | 1.609% | 1.796% | 0.6906% | 1.937% | 3.336% | 2.214% | 1.115% | 1.457% | 4.13% | 1.726% | 3.355%
| PDM3 |  | -2.793% | 9.171% | -21.1% | -9.407% | -17.12% | 3.467% | 9.304% | 1.859% | 28.61% | -22.2% | 21.3%
| Z-Score (Absolute) |  | 0.28 | 0.44 | 1.37 | 0.67 | 1.2 | 0.1 | 0.45 | 0 | 1.6 | 1.44 | 1.16
| Analytical results for Cs in OREAS 462 (Certified Value 0.45 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 0.44 | 0.45 | 0.49 | 0.45 | 0.42 | 0.43172 | 0.4 | < 0.5 | 0.49 | 0.36 | 0.53733 | < 0.5 | 0.4 |  | 1
| 1 | 2 | 0.45 | 0.44 | 0.46 | 0.45 | 0.45 | 0.42379 | 0.5 | 0.5 | 0.48 | 0.35 | 0.55311 | < 0.5 | 0.4 |  | 0
| 1 | 3 | 0.28 | 0.48 | 0.51 | 0.44 | 0.42 | 0.41389 | 0.5 | < 0.5 | 0.48 | 0.36 | 0.59074 | < 0.5 | 0.4 |  | 16
| 1 | 4 | 0.37 | 0.47 | 0.49 | 0.45 | 0.41 | 0.42845 | 0.5 | 0.5 | 0.48 | 0.34 | 0.53597 | 0.5 | 0.4 |  | 0.45029
| 1 | 5 | 0.32 | 0.45 | 0.5 | 0.45 | 0.41 | 0.39727 | 0.5 | < 0.5 | 0.48 | 0.34 | 0.5691 | 0.5 | 0.4 |  | 77
| 1 | 6 | 0.29 | 0.44 | 0.49 | 0.44 | 0.42 | 0.4168 | 0.5 | < 0.5 | 0.49 | 0.35 | 0.52853 | < 0.5 | 0.4
| Mean |  | 0.35833 | 0.455 | 0.49 | 0.44667 | 0.42167 | 0.41865 | 0.48333 | 0.5 | 0.48333 | 0.35 | 0.55246 | 0.5 | 0.4
| Median |  | 0.345 | 0.45 | 0.49 | 0.45 | 0.42 | 0.4203 | 0.5 | 0.5 | 0.48 | 0.35 | 0.54522 | 0.5 | 0.4
| Std Dev. |  | 0.07414 | 0.016432 | 0.016733 | 0.005164 | 0.01472 | 0.01246 | 0.040825 | 0 | 0.005164 | 0.0089443 | 0.023731 | 0 | 6.0809e-17
| Rel.Std.Dev. |  | 20.69% | 3.611% | 3.415% | 1.156% | 3.491% | 2.976% | 8.447% | 0% | 1.068% | 2.556% | 4.295% | 0% | 1.52e-14%
| PDM3 |  | -20.42% | 1.046% | 8.819% | -0.8049% | -6.357% | -7.026% | 7.338% | 11.04% | 7.338% | -22.27% | 22.69% | 11.04% | -11.17%
| Z-Score (Absolute) |  | 1.67 | 0 | 0.61 | 0.14 | 0.67 | 0.63 | 0.49 | 0.78 | 0.49 | 1.82 | 1.69 | 0.78 | 0.95
| Analytical results for Cu in OREAS 462 (Certified Value 61 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 60.7 | 61.8 | 63.2 | 59 | 58.1 | 59.4503 | 63.8 | 82 | 63.2 | 98 | 60.322 | 31.5 | 60 | 82 | 1
| 1 | 2 | 60.3 | 61.3 | 61 | 58.9 | 57.9 | 61.097 | 63.2 | 74 | 62.1 | 98 | 60.463 | 31 | 60 | 78 | 19
| 1 | 3 | 62 | 67.7 | 65.8 | 59 | 57.8 | 59.152 | 63.1 | 76 | 63.9 | 96 | 60.6959 | 30.9 | 60 | 78 | 16
| 1 | 4 | 60.1 | 61.5 | 63.4 | 58.8 | 54.1 | 61.4742 | 62.9 | 62 | 63.4 | 99 | 61.179 | 31.6 | 60 | 80 | 60.8857
| 1 | 5 | 60.2 | 61.5 | 64.6 | 57.6 | 59.7 | 60.7865 | 62.5 | 71 | 63.2 | 103 | 60.7336 | 32.7 | 60 | 86 | 78
| 1 | 6 | 60.3 | 62.5 | 60.3 | 59 | 55.5 | 59.3948 | 62.6 | 75 | 62.4 | 104 | 60.7721 | 31.4 | 60 | 78
| Mean |  | 60.6 | 62.7167 | 63.05 | 58.7167 | 57.1833 | 60.2258 | 63.0167 | 73.3333 | 63.0333 | 99.6667 | 60.6943 | 31.5167 | 60 | 80.3333
| Median |  | 60.3 | 61.65 | 63.3 | 58.95 | 57.85 | 60.1184 | 63 | 74.5 | 63.2 | 98.5 | 60.7148 | 31.45 | 60 | 79
| Std Dev. |  | 0.71554 | 2.4774 | 2.0917 | 0.55287 | 2.0203 | 1.0077 | 0.47081 | 6.6232 | 0.66533 | 3.1411 | 0.29484 | 0.64317 | 0 | 3.2042
| Rel.Std.Dev. |  | 1.181% | 3.95% | 3.317% | 0.9416% | 3.533% | 1.673% | 0.7471% | 9.032% | 1.056% | 3.152% | 0.4858% | 2.041% | 0% | 3.989%
| PDM3 |  | -0.4692% | 3.007% | 3.555% | -3.562% | -6.081% | -1.084% | 3.5% | 20.44% | 3.527% | 63.69% | -0.3144% | -48.24% | -1.455% | 31.94%
| Z-Score (Absolute) |  | 0.22 | 0.19 | 0.68 | 0.92 | 1.25 | 0.36 | 0.67 | 4.46 | 0.67 | 14.14 | 0.19 | 11 | 0.44 | 7.03
| Analytical results for Dy in OREAS 462 (Certified Value 51 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 52.5 | 49.6 | 52.9 | 50.4 | 50.6 | 52.1037 | 51 | 50.9 | 51.7 | 48.95 | 52.9055 | 51.1 | 47.6 |  | 1
| 1 | 2 | 53.4 | 51.1 | 52.1 | 50.7 | 54 | 52.0391 | 49.9 | 51 | 52 | 50.85 | 53.7253 | 50.3 | 49.4 |  | 5
| 1 | 3 | 48.3 | 55 | 55.3 | 50.4 | 52.1 | 51.3669 | 50.7 | 50.6 | 51.9 | 48.56 | 53.6278 | 50.2 | 50.2 |  | 16
| 1 | 4 | 51.4 | 49.7 | 52.8 | 50.8 | 49.3 | 51.3395 | 49.7 | 51.4 | 51.9 | 49.58 | 52.5691 | 50.6 | 48.9 |  | 50.5936
| 1 | 5 | 50 | 49.9 | 53.6 | 48.4 | 50.2 | 51.7658 | 50.4 | 51.3 | 51.9 | 48.72 | 54.5901 | 49.6 | 49 |  | 79
| 1 | 6 | 47.6 | 52.2 | 51.4 | 51 | 50.6 | 52.0226 | 50.3 | 51.8 | 52.5 | 50.06 | 53.2776 | 50.5 | 50
| Mean |  | 50.5333 | 51.25 | 53.0167 | 50.2833 | 51.1333 | 51.7729 | 50.3333 | 51.1667 | 51.9833 | 49.4533 | 53.4492 | 50.3833 | 49.1833
| Median |  | 50.7 | 50.5 | 52.85 | 50.55 | 50.6 | 51.8942 | 50.35 | 51.15 | 51.9 | 49.265 | 53.4527 | 50.4 | 49.2
| Std Dev. |  | 2.3114 | 2.0945 | 1.3467 | 0.95167 | 1.6705 | 0.34506 | 0.48442 | 0.42269 | 0.27142 | 0.88588 | 0.70843 | 0.49565 | 0.9347
| Rel.Std.Dev. |  | 4.574% | 4.087% | 2.54% | 1.893% | 3.267% | 0.6665% | 0.9624% | 0.8261% | 0.5221% | 1.791% | 1.325% | 0.9838% | 1.9%
| PDM3 |  | -0.1191% | 1.297% | 4.789% | -0.6133% | 1.067% | 2.331% | -0.5144% | 1.133% | 2.747% | -2.254% | 5.644% | -0.4156% | -2.787%
| Z-Score (Absolute) |  | 0.03 | 0.07 | 2.73 | 0.11 | 0 | 1.35 | 0.25 | 0.67 | 1.58 | 1.23 | 3.21 | 0.2 | 1.53
| Analytical results for Er in OREAS 462 (Certified Value 10.8 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 11.55 | 11 | 14.1 | 10.85 | 10.35 | 11.1397 | 11 | 10.6 | 16.6 | 10.13 | 18.3174 | 10.85 | 11 |  | 1
| 1 | 2 | 11.85 | 11.4 | 13.55 | 10.85 | 11.2 | 10.3216 | 11.1 | 10.5 | 16.5 | 10.21 | 19.0377 | 10.7 | 11.2 |  | 6
| 1 | 3 | 10.75 | 12 | 14.35 | 10.8 | 10.9 | 10.6738 | 10.7 | 10.8 | 16.6 | 10.23 | 19.1753 | 10.75 | 11.6 |  | 16
| 1 | 4 | 11.15 | 11.1 | 13.65 | 10.75 | 10.4 | 10.3637 | 11 | 10.5 | 16.7 | 10.07 | 18.7964 | 10.9 | 11.3 |  | 10.8497
| 1 | 5 | 11.15 | 10.95 | 14.2 | 10.3 | 10.5 | 11.16 | 10.9 | 10.5 | 16.5 | 10.37 | 18.741 | 10.95 | 11.2 |  | 80
| 1 | 6 | 10.95 | 11.6 | 13.25 | 11.1 | 10.5 | 10.6446 | 10.7 | 10.9 | 16.7 | 10.17 | 17.2294 | 10.7 | 11.3
| Mean |  | 11.2333 | 11.3417 | 13.85 | 10.775 | 10.6417 | 10.7172 | 10.9 | 10.6333 | 16.6 | 10.1967 | 18.5495 | 10.8083 | 11.2667
| Median |  | 11.15 | 11.25 | 13.875 | 10.825 | 10.5 | 10.6592 | 10.95 | 10.55 | 16.6 | 10.19 | 18.7687 | 10.8 | 11.25
| Std Dev. |  | 0.40208 | 0.40794 | 0.43012 | 0.2622 | 0.33529 | 0.36419 | 0.16733 | 0.17512 | 0.089443 | 0.1025 | 0.7105 | 0.10685 | 0.19664
| Rel.Std.Dev. |  | 3.579% | 3.597% | 3.106% | 2.433% | 3.151% | 3.398% | 1.535% | 1.647% | 0.5388% | 1.005% | 3.83% | 0.9886% | 1.745%
| PDM3 |  | 3.536% | 4.534% | 27.65% | -0.6887% | -1.918% | -1.221% | 0.4634% | -1.994% | 53% | -6.019% | 70.97% | -0.3815% | 3.843%
| Z-Score (Absolute) |  | 0.61 | 0.81 | 5.43 | 0.06 | 0.68 | 0.34 | 0 | 0.49 | 10.48 | 1.29 | 14.55 | 0.17 | 0.67
| Analytical results for Eu in OREAS 462 (Certified Value 80 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 89.8 | 77.9 | 87 | 84.9 | 80.3 | 77.9981 | 80.1 | 83.5 | 73.4 | 72.65 | 93.4423 | 77.4 | 74.1 |  | 1
| 1 | 2 | 91.4 | 80.4 | 83.9 | 84.7 | 86.2 | 75.3285 | 79 | 82.8 | 74.4 | 75.82 | 93.5272 | 78.1 | 77.4 |  | 7
| 1 | 3 | 83 | 89.4 | 90 | 84.2 | 82.4 | 76.28 | 78.7 | 81.3 | 74.5 | 74.9 | 92.5942 | 77.7 | 77 |  | 16
| 1 | 4 | 88.6 | 81.8 | 84.7 | 83.7 | 78.2 | 76.7095 | 78.3 | 82.2 | 74.7 | 74.89 | 92.6741 | 78.9 | 76.5 |  | 79.8818
| 1 | 5 | 84.1 | 79.6 | 86.1 | 81.3 | 81 | 78.9528 | 79.6 | 84 | 73.4 | 74.33 | 93.8241 | 77.3 | 75.3 |  | 81
| 1 | 6 | 81 | 80.8 | 81.9 | 85.3 | 81 | 77.2377 | 78.4 | 82.5 | 74.8 | 73.45 | 94.5748 | 77.9 | 76.8
| Mean |  | 86.3167 | 81.65 | 85.6 | 84.0167 | 81.5167 | 77.0844 | 79.0167 | 82.7167 | 74.2 | 74.34 | 93.4395 | 77.8833 | 76.1833
| Median |  | 86.35 | 80.6 | 85.4 | 84.45 | 81 | 76.9736 | 78.85 | 82.65 | 74.45 | 74.61 | 93.4847 | 77.8 | 76.65
| Std Dev. |  | 4.1802 | 4.0158 | 2.7914 | 1.4428 | 2.6731 | 1.2822 | 0.70828 | 0.95795 | 0.63561 | 1.1362 | 0.74092 | 0.58109 | 1.2449
| Rel.Std.Dev. |  | 4.843% | 4.918% | 3.261% | 1.717% | 3.279% | 1.663% | 0.8964% | 1.158% | 0.8566% | 1.528% | 0.7929% | 0.7461% | 1.634%
| PDM3 |  | 8.056% | 2.214% | 7.158% | 5.176% | 2.047% | -3.502% | -1.083% | 3.549% | -7.113% | -6.937% | 16.97% | -2.502% | -4.63%
| Z-Score (Absolute) |  | 1.07 | 0 | 0.95 | 0.77 | 0.08 | 0.52 | 0.19 | 0.45 | 1.02 | 0.99 | 2.3 | 0.38 | 0.67
| Analytical results for Fe in OREAS 462 (Certified Value 33.05 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 29.9 | 32.9 | 31.7 | 34.5 | 30.8 | 34.4216 | 32.28 | 33.49 | 33 | 33.01 | 32.479 | - | 33.4255 | 33.9 | 1
| 1 | 2 | 30.4 | 32.9 | 31.6 | 34.7 | 32.2 | 34.7457 | 31.95 | 33.08 | 33.3 | 34.04 | 32.9651 | > 15 | 33.193 | 33.6 | 0
| 1 | 3 | 29.5 | 33.5 | 31.7 | 34.4 | 33 | 34.67 | 31.99 | 33.56 | 33.1 | 32.57 | 32.6797 | > 15 | 33.105 | 33.6 | 16
| 1 | 4 | 29.3 | 32.4 | 32.1 | 34.6 | 30.5 | 34.5351 | 32.34 | 33.44 | 33.3 | 33.04 | 32.825 | > 15 | 33.1235 | 34.1 | 33.0543
| 1 | 5 | 31.1 | 32.5 | 31.7 | 33.5 | 37.8 | 34.758 | 32.25 | 33.06 | 32.7 | 33.84 | 32.3427 | > 15 | 33.05 | 34.5 | 82
| 1 | 6 | 28.7 | 32.9 | 32.1 | 34.9 | 31.1 | 34.4562 | 32.07 | 32.97 | 32.8 | 32.44 | 32.493 | > 15 | 32.8795 | 33.6
| Mean |  | 29.8167 | 32.85 | 31.8167 | 34.4333 | 32.5667 | 34.5978 | 32.1467 | 33.2667 | 33.0333 | 33.1567 | 32.6307 |  | 33.1294 | 33.8833
| Median |  | 29.7 | 32.9 | 31.7 | 34.55 | 31.65 | 34.6025 | 32.16 | 33.26 | 33.05 | 33.025 | 32.5864 |  | 33.1142 | 33.75
| Std Dev. |  | 0.84951 | 0.38859 | 0.22286 | 0.48854 | 2.7296 | 0.14683 | 0.16428 | 0.2575 | 0.25033 | 0.65424 | 0.23533 |  | 0.1795 | 0.3656
| Rel.Std.Dev. |  | 2.849% | 1.183% | 0.7005% | 1.419% | 8.382% | 0.4244% | 0.511% | 0.7741% | 0.7578% | 1.973% | 0.7212% |  | 0.5418% | 1.079%
| PDM3 |  | -9.795% | -0.618% | -3.744% | 4.172% | -1.475% | 4.67% | -2.746% | 0.6426% | -0.06335% | 0.3098% | -1.281% |  | 0.2273% | 2.508%
| Z-Score (Absolute) |  | 2.55 | 0.15 | 0.97 | 1.26 | 1.2 | 1.24 | 0.7 | 0.19 | 0 | 0.1 | 0.32 | NA | 0.08 | 0.67
| Analytical results for Ga in OREAS 462 (Certified Value 58 ppm) | [15 empty] | Certified
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 10; 08: 12; 09: 22 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS | 0
0
| 1 | 1 | 01: 53.3; 02: 54.8; 03: 54.6; 05: 71.7; 06: 34.7037; 07: < 0.05; 08: 54.7053; 09: 33.2 | 1
| 1 | 2 | 01: 55.2; 02: 52.9; 03: 53.8; 05: 75.1; 06: 34.4022; 07: < 0.05; 08: 55.9504; 09: 33 | 0
| 1 | 3 | 01: 60.3; 02: 56.8; 03: 54.2; 05: 72.5; 06: 33.7991; 07: < 0.05; 08: 56.0375; 09: 30.6 | 16
| 1 | 4 | 01: 55.3; 02: 55; 03: 53.7; 05: 70.5; 06: 35.8769; 07: < 0.05; 08: 54.9919; 09: 32.2 | 58.3472
| 1 | 5 | 01: 55.1; 02: 56; 03: 52.7; 05: 70.7; 06: 34.6183; 07: < 0.05; 08: 55.8818; 09: 32.6 | 83
| 1 | 6 | 01: 56.2; 02: 52.5; 03: 54.2; 05: 71.1; 06: 36.8409; 07: < 0.05; 08: 54.6491; 09: 32.6
| Mean |  | 01: 55.9; 02: 54.6667; 03: 53.8667; 05: 71.9333; 06: 35.0402; 08: 55.3693; 09: 32.3667
| Median |  | 01: 55.25; 02: 54.9; 03: 54; 05: 71.4; 06: 34.661; 08: 55.4368; 09: 32.6
| Std Dev. |  | 01: 2.3537; 02: 1.6896; 03: 0.65625; 05: 1.7131; 06: 1.1119; 08: 0.65556; 09: 0.9331
| Rel.Std.Dev. |  | 01: 4.211%; 02: 3.091%; 03: 1.218%; 05: 2.381%; 06: 3.173%; 08: 1.184%; 09: 2.883%
| PDM3 |  | 01: -4.194%; 02: -6.308%; 03: -7.679%; 05: 23.28%; 06: -39.95%; 08: -5.104%; 09: -44.53%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0; 03: 0.44; 05: 9.44; 06: 10.73; 07: NA; 08: 0.38; 09: 12.19
| Analytical results for Gd in OREAS 462 (Certified Value 164 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 169.5 | 160.5 | 199 | 154.5 | 166 | 166.8941 | 169.2 | 169.4 | 85.6 | 160.53 | 243.0774 | 164.7 | 160 |  | 1
| 1 | 2 | 171 | 166.5 | 194 | 155 | 180.5 | 175.3779 | 167.6 | 165.4 | 85.7 | 160.19 | 242.6132 | 164.9 | 169 |  | 8
| 1 | 3 | 153 | 183.5 | 203 | 155.5 | 171 | 173.334 | 165.3 | 168 | 83.5 | 157.02 | 241.5687 | 162.7 | 164 |  | 16
| 1 | 4 | 165 | 160.5 | 189 | 152.5 | 162.5 | 171.8094 | 167.8 | 165.4 | 89.7 | 158.66 | 242.8719 | 164.7 | 164 |  | 163.8761
| 1 | 5 | 158 | 166.5 | 191.5 | 148 | 168.5 | 164.9212 | 167.9 | 168 | 90.8 | 156.26 | 243.4772 | 161.9 | 163 |  | 84
| 1 | 6 | 149.5 | 168 | 179.5 | 155.5 | 166 | 167.3066 | 165.9 | 167.1 | 89.1 | 160.06 | 242.9835 | 166.5 | 167
| Mean |  | 161 | 167.5833 | 192.6667 | 153.5 | 169.0833 | 169.9405 | 167.2833 | 167.2167 | 87.4 | 158.7867 | 242.7653 | 164.2333 | 164.5
| Median |  | 161.5 | 166.5 | 192.75 | 154.75 | 167.25 | 169.558 | 167.7 | 167.55 | 87.4 | 159.36 | 242.9277 | 164.7 | 164
| Std Dev. |  | 8.8713 | 8.4405 | 8.2077 | 2.9155 | 6.2723 | 4.1469 | 1.4331 | 1.588 | 2.8664 | 1.7973 | 0.65087 | 1.6621 | 3.1464
| Rel.Std.Dev. |  | 5.51% | 5.037% | 4.26% | 1.899% | 3.71% | 2.44% | 0.8567% | 0.9497% | 3.28% | 1.132% | 0.2681% | 1.012% | 1.913%
| PDM3 |  | -1.755% | 2.262% | 17.57% | -6.332% | 3.178% | 3.701% | 2.079% | 2.038% | -46.67% | -3.106% | 48.14% | 0.218% | 0.3807%
| Z-Score (Absolute) |  | 0.67 | 0.02 | 5.43 | 1.91 | 0.44 | 1.05 | 0.54 | 0.52 | 14.85 | 1.1 | 15.08 | 0.05 | 0
| Analytical results for Ge in OREAS 462 (Indicative Value 6 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | [5 empty] | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | [5 empty] | 2
2
| 1 | 1 | 3.65 | 5.39 | 5.43 | 4.26 | 7.5 | 6 | 9.5 | < 0.05 | 6.31 | [5 empty] | 1
| 1 | 2 | 4.01 | 5.75 | 5.25 | 4.17 | 7.59 | 5.9 | 9 | < 0.05 | 6.89 | [5 empty] | 21
| 1 | 3 | 3.51 | 5.69 | 5.48 | 4.04 | 7.61 | 6.1 | 8.7 | < 0.05 | 6.63 | [5 empty] | 16
| 1 | 4 | 3.52 | 5.68 | 5.38 | 4.25 | 7.19 | 6.1 | 9.4 | < 0.05 | 6.5 | [5 empty] | 6
| 1 | 5 | 3.5 | 5.53 | 5.4 | 4.06 | 7.5 | 5.9 | 9.3 | < 0.05 | 6.4 | [5 empty] | 27
| 1 | 6 | 3.37 | 5.69 | 4.97 | 4.3 | 7.43 | 5.9 | 10.4 | < 0.05 | 6.47
| Mean |  | 3.5933 | 5.6217 | 5.3183 | 4.18 | 7.47 | 5.9833 | 9.3833 |  | 6.5333
| Median |  | 3.515 | 5.685 | 5.39 | 4.21 | 7.5 | 5.95 | 9.35 |  | 6.485
| Std Dev. |  | 0.22259 | 0.13512 | 0.18713 | 0.10936 | 0.15218 | 0.098319 | 0.57764 |  | 0.20461
| Rel.Std.Dev. |  | 6.195% | 2.404% | 3.519% | 2.616% | 2.037% | 1.643% | 6.156% |  | 3.132%
| PDM3 |  | -40.11% | -6.306% | -11.36% | -30.33% | 24.5% | -0.2778% | 56.39% |  | 8.889%
| Z-Score (Absolute) |  | 1.31 | 0.1 | 0.28 | 0.93 | 0.96 | 0.1 | 2.05 | NA | 0.42
| Analytical results for Hf in OREAS 462 (Certified Value 7.4 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 8.7 | 6.6 | 8.4 | 7.8 | 7.7 | 6.5827 | 6.3 | 7.31 | 7.98 | 6.4 |  |  |  |  | 1
| 1 | 2 | 8.9 | 7 | 8.6 | 7.7 | 8 | 6.307 | 6 | 7.2 | 7.46 | 6.6 |  |  |  |  | 9
| 1 | 3 | 8.3 | 7.3 | 8.7 | 8 | 7.8 | 6.3619 | 6.1 | 7.1 | 7.69 | 6.6 |  |  |  |  | 16
| 1 | 4 | 8.5 | 6.8 | 9 | 7.9 | 7.5 | 6.4226 | 6.4 | 7.09 | 7.78 | 6.4 |  |  |  |  | 7.3979
| 1 | 5 | 8.3 | 7.1 | 8.9 | 7.6 | 7.6 | 6.54 | 6.3 | 6.92 | 7.75 | 6.6 |  |  |  |  | 85
| 1 | 6 | 8.4 | 7.2 | 8.8 | 7.8 | 7.8 | 6.6198 | 6.5 | 7.07 | 7.99 | 6.8
| Mean |  | 8.5167 | 7 | 8.7333 | 7.8 | 7.7333 | 6.4723 | 6.2667 | 7.115 | 7.775 | 6.5667
| Median |  | 8.45 | 7.05 | 8.75 | 7.8 | 7.75 | 6.4813 | 6.3 | 7.095 | 7.765 | 6.6
| Std Dev. |  | 0.24014 | 0.26077 | 0.21602 | 0.14142 | 0.17512 | 0.12691 | 0.18619 | 0.13126 | 0.19766 | 0.15055
| Rel.Std.Dev. |  | 2.82% | 3.725% | 2.474% | 1.813% | 2.264% | 1.961% | 2.971% | 1.845% | 2.542% | 2.293%
| PDM3 |  | 15.12% | -5.379% | 18.05% | 5.435% | 4.534% | -12.51% | -15.29% | -3.824% | 5.097% | -11.24%
| Z-Score (Absolute) |  | 1.15 | 0.45 | 1.38 | 0.4 | 0.33 | 1 | 1.22 | 0.33 | 0.37 | 0.9
| Analytical results for Ho in OREAS 462 (Certified Value 6.18 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 6.77 | 6.06 | 6.86 | 6.1 | 5.9 | 6.202 | 6.2 | 5.9 | 7.14 | 5.8 | 7.3256 | 6.1 | 5.58 |  | 1
| 1 | 2 | 6.97 | 6.18 | 6.75 | 5.97 | 6.68 | 6.3356 | 6.2 | 6.1 | 6.9 | 5.96 | 7.6744 | 6.05 | 5.82 |  | 10
| 1 | 3 | 6.25 | 6.66 | 7.24 | 6.07 | 6.54 | 6.3329 | 6.3 | 6 | 7.59 | 5.74 | 7.3825 | 6.05 | 5.88 |  | 16
| 1 | 4 | 6.59 | 6.04 | 6.81 | 6.01 | 5.71 | 6.1513 | 6.2 | 6 | 7.88 | 5.78 | 7.2439 | 6.15 | 5.76 |  | 6.1793
| 1 | 5 | 6.43 | 5.95 | 7.05 | 5.77 | 6.36 | 6.2007 | 6.2 | 6.1 | 7.32 | 5.81 | 7.7806 | 6.1 | 5.8 |  | 86
| 1 | 6 | 6.13 | 6.35 | 6.72 | 6.13 | 5.97 | 6.2234 | 6.1 | 6 | 6.82 | 5.81 | 7.5079 | 6.2 | 5.84
| Mean |  | 6.5233 | 6.2067 | 6.905 | 6.0083 | 6.1933 | 6.241 | 6.2 | 6.0167 | 7.275 | 5.8167 | 7.4858 | 6.1083 | 5.78
| Median |  | 6.51 | 6.12 | 6.835 | 6.04 | 6.165 | 6.2127 | 6.2 | 6 | 7.23 | 5.805 | 7.4452 | 6.1 | 5.81
| Std Dev. |  | 0.31715 | 0.26151 | 0.20127 | 0.1306 | 0.38842 | 0.076023 | 0.063246 | 0.075277 | 0.4082 | 0.075011 | 0.20871 | 0.058452 | 0.10583
| Rel.Std.Dev. |  | 4.862% | 4.213% | 2.915% | 2.174% | 6.272% | 1.218% | 1.02% | 1.251% | 5.611% | 1.29% | 2.788% | 0.9569% | 1.831%
| PDM3 |  | 5.568% | 0.4429% | 11.74% | -2.767% | 0.2271% | 0.9979% | 0.335% | -2.632% | 17.73% | -5.869% | 21.14% | -1.148% | -6.462%
| Z-Score (Absolute) |  | 1.14 | 0.02 | 2.24 | 0.67 | 0.02 | 0.14 | 0 | 0.64 | 3.78 | 1.35 | 4.52 | 0.32 | 1.34
| Analytical results for In in OREAS 462 (Certified Value 0.75 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 0.751 | 0.723 | 0.914 | 0.777 | 0.758 | 0.70439 | 0.8 | 0.75 | 0.758 | 0.72 | 0.8 | 0.66 |  |  | 1
| 1 | 2 | 0.754 | 0.739 | 0.87 | 0.755 | 0.813 | 0.72349 | 0.8 | 0.77 | 0.76 | 0.73 | 0.8 | 0.7 |  |  | 16
| 1 | 3 | 0.754 | 0.797 | 0.922 | 0.787 | 0.794 | 0.71492 | 0.8 | 0.89 | 0.753 | 0.74 | 0.7 | 0.68 |  |  | 16
| 1 | 4 | 0.723 | 0.753 | 0.931 | 0.739 | 0.764 | 0.72156 | 0.8 | 0.83 | 0.756 | 0.73 | 0.8 | 0.66 |  |  | 0.74891
| 1 | 5 | 0.746 | 0.741 | 0.912 | 0.76 | 0.778 | 0.73332 | 0.8 | 0.73 | 0.749 | 0.73 | 0.7 | 0.68 |  |  | 87
| 1 | 6 | 0.756 | 0.762 | 0.871 | 0.764 | 0.781 | 0.69887 | 0.8 | 0.73 | 0.745 | 0.73 | 0.7 | 0.68
| Mean |  | 0.74733 | 0.7525 | 0.90333 | 0.76367 | 0.78133 | 0.71609 | 0.8 | 0.78333 | 0.7535 | 0.73 | 0.75 | 0.67667
| Median |  | 0.7525 | 0.747 | 0.913 | 0.762 | 0.7795 | 0.71824 | 0.8 | 0.76 | 0.7545 | 0.73 | 0.75 | 0.68
| Std Dev. |  | 0.01242 | 0.025517 | 0.026303 | 0.016848 | 0.020096 | 0.012779 | 1.2162e-16 | 0.064083 | 0.0056833 | 0.0063246 | 0.054772 | 0.015055
| Rel.Std.Dev. |  | 1.662% | 3.391% | 2.912% | 2.206% | 2.572% | 1.784% | 1.52e-14% | 8.181% | 0.7543% | 0.8664% | 7.303% | 2.225%
| PDM3 |  | -0.2099% | 0.48% | 20.62% | 1.971% | 4.33% | -4.381% | 6.823% | 4.597% | 0.6135% | -2.524% | 0.1462% | -9.646%
| Z-Score (Absolute) |  | 0.03 | 0.02 | 6.02 | 0.43 | 1.13 | 1.48 | 1.88 | 0.36 | 0.02 | 0.92 | 0.12 | 3.06
| Analytical results for K in OREAS 462 (Certified Value 0.118 wt.%) | [15 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.11 | 0.12 | 0.12 | 0.12 | 0.11 | 0.12 | 0.13 | 0.128 | 0.115 | 0.12 | 0.11 | 0.1 |  |  | 1
| 1 | 2 | 0.11 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.11 | 0.129 | 0.1188 | 0.12 | 0.11 | 0.1 |  |  | 11
| 1 | 3 | 0.09 | 0.12 | 0.12 | 0.11 | 0.12 | 0.12 | 0.12 | 0.128 | 0.1149 | 0.12 | 0.11 | 0.1 |  |  | 16
| 1 | 4 | 0.1 | 0.12 | 0.12 | 0.12 | 0.11 | 0.12 | 0.12 | 0.13 | 0.1165 | 0.12 | 0.11 | 0.1 |  |  | 0.11843
| 1 | 5 | 0.1 | 0.12 | 0.12 | 0.11 | 0.14 | 0.12 | 0.12 | 0.126 | 0.114 | 0.12 | 0.11 | 0.11 |  |  | 88
| 1 | 6 | 0.09 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.11 | 0.126 | 0.1138 | 0.12 | 0.11 | 0.1
| Mean |  | 0.1 | 0.12 | 0.12 | 0.11667 | 0.12 | 0.12 | 0.11833 | 0.12783 | 0.1155 | 0.12 | 0.11 | 0.10167
| Median |  | 0.1 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.12 | 0.128 | 0.11495 | 0.12 | 0.11 | 0.1
| Std Dev. |  | 0.0089443 | 0 | 0 | 0.005164 | 0.010954 | 0 | 0.0075277 | 0.0016021 | 0.0018783 | 0 | 0 | 0.0040825
| Rel.Std.Dev. |  | 8.944% | 0% | 0% | 4.426% | 9.129% | 0% | 6.361% | 1.253% | 1.626% | 0% | 0% | 4.016%
| PDM3 |  | -15.56% | 1.323% | 1.323% | -1.492% | 1.323% | 1.323% | -0.08444% | 7.937% | -2.477% | 1.323% | -7.121% | -14.16%
| Z-Score (Absolute) |  | 4.72 | 0.67 | 0.67 | 0.22 | 0.4 | 0.67 | 0.22 | 2.79 | 0.54 | 0.67 | 2.02 | 4.27
| Analytical results for La in OREAS 462 (Certified Value 3669 ppm) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 3270 | 3710 | 2680 | 2760 | 2530 | 3961.9659 | 3582.8 | 3575.6 | 3790 | 3510.14 | 3883.7076 | 3821 | 3610 |  | 1
| 1 | 2 | 3320 | 3690 | 2700 | 2840 | 2560 | 3938.0074 | 3553.5 | 3576 | 3790 | 3620.31 | 3844.7882 | 3793 | 3730 |  | 0
| 1 | 3 | 2860 | 3780 | 2720 | 2850 | 2580 | 3886.7358 | 3568.1 | 3569.5 | 3850 | 3515.64 | 3829.0546 | 3818 | 3750 |  | 16
| 1 | 4 | 3070 | 3640 | 2740 | 2830 | 2540 | 3881.6968 | 3547.7 | 3619.1 | 3870 | 3530.42 | 3751.8977 | 3879 | 3670 |  | 3669.026
| 1 | 5 | 3080 | 3630 | 2720 | 2770 | 2860 | 3823.4323 | 3546.9 | 3659.4 | 3800 | 3518.96 | 3890.3817 | 3809.5 | 3740 |  | 89
| 1 | 6 | 2880 | 3690 | 2750 | 2800 | 2590 | 3875.3633 | 3506.9 | 3595.3 | 3850 | 3517.31 | 3888.9489 | 3883.5 | 3750
| Mean |  | 3080 | 3690 | 2718.3333 | 2808.3333 | 2610 | 3894.5336 | 3550.9833 | 3599.15 | 3825 | 3535.4633 | 3848.1298 | 3834 | 3708.3333
| Median |  | 3075 | 3690 | 2720 | 2815 | 2570 | 3884.2163 | 3550.6 | 3585.65 | 3825 | 3518.135 | 3864.2479 | 3819.5 | 3735
| Std Dev. |  | 190.8926 | 54.037 | 25.6255 | 37.6386 | 124.5793 | 49.1642 | 25.6219 | 34.6772 | 35.6371 | 42.0964 | 53.5777 | 37.8985 | 56.7157
| Rel.Std.Dev. |  | 6.198% | 1.464% | 0.9427% | 1.34% | 4.773% | 1.262% | 0.7215% | 0.9635% | 0.9317% | 1.191% | 1.392% | 0.9885% | 1.529%
| PDM3 |  | -16.05% | 0.5717% | -25.91% | -23.46% | -28.86% | 6.146% | -3.217% | -1.904% | 4.251% | -3.64% | 4.882% | 4.496% | 1.071%
| Z-Score (Absolute) |  | 1.19 | 0.26 | 2.53 | 2.27 | 2.98 | 0.85 | 0.14 | 0 | 0.65 | 0.18 | 0.71 | 0.67 | 0.37
| Analytical results for Li in OREAS 462 (Certified Value 10.8 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 1
1
| 1 | 1 | 9.3 | 10.4 | 10.7 | 11.6 | 8.9 | 11.2504 | 11 | 12 | 10.9 | 8.3 | 12.9671 | 12 | 10 | 11 | 1
| 1 | 2 | 9.3 | 10.6 | 10 | 11.5 | 9.5 | 12.4818 | 10.7 | 12 | 10.9 | 8.7 | 13.3408 | 12 | 10 | 11 | 0
| 1 | 3 | 9 | 11.6 | 11.3 | 11.5 | 8.7 | 12.1072 | 10.6 | 12 | 10.9 | 8.7 | 13.0247 | 12 | 10 | 11 | 16
| 1 | 4 | 9.1 | 10.3 | 10.6 | 11.4 | 8.7 | 11.9141 | 10.5 | 13 | 11 | 7.6 | 12.8734 | 12 | 15 | 11 | 10.8383
| 1 | 5 | 8.8 | 10.8 | 11.2 | 11.1 | 8.4 | 11.5798 | 10.6 | 12 | 10.8 | 8.2 | 13.0994 | 12 | 10 | 11 | 90
| 1 | 6 | 8.8 | 10.8 | 10.7 | 11.6 | 9 | 12.2792 | 10.6 | 12 | 11 | 8.7 | 12.6032 | 12 | 10 | 11
| Mean |  | 9.05 | 10.75 | 10.75 | 11.45 | 8.8667 | 11.9354 | 10.6667 | 12.1667 | 10.9167 | 8.3667 | 12.9848 | 12 | 10.8333 | 11
| Median |  | 9.05 | 10.7 | 10.7 | 11.5 | 8.8 | 12.0107 | 10.6 | 12 | 10.9 | 8.5 | 12.9959 | 12 | 10 | 11
| Std Dev. |  | 0.22583 | 0.46368 | 0.46797 | 0.18708 | 0.37238 | 0.45657 | 0.17512 | 0.40825 | 0.075277 | 0.43665 | 0.24486 | 0 | 2.0412 | 0
| Rel.Std.Dev. |  | 2.495% | 4.313% | 4.353% | 1.634% | 4.2% | 3.825% | 1.642% | 3.355% | 0.6896% | 5.219% | 1.886% | 0% | 18.84% | 0%
| PDM3 |  | -16.5% | -0.8151% | -0.8151% | 5.643% | -18.19% | 10.12% | -1.584% | 12.26% | 0.7226% | -22.8% | 19.8% | 10.72% | -0.04627% | 1.491%
| Z-Score (Absolute) |  | 1.5 | 0.1 | 0.1 | 0.47 | 1.66 | 0.87 | 0.17 | 1.07 | 0.03 | 2.07 | 1.74 | 0.93 | 0.03 | 0.1
| Analytical results for Lu in OREAS 462 (Certified Value 0.53 ppm) | [15 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 0.74 | 0.55 | 0.6 | 0.5 | 0.54 | 0.43276 | 0.5 | 0.4 | 0.651 | 0.41 | 0.59907 | 0.5 | 0.5 |  | 1
| 1 | 2 | 0.76 | 0.57 | 0.56 | 0.49 | 0.57 | 0.4544 | 0.5 | 0.36 | 0.638 | 0.41 | 0.62588 | 0.5 | 0.52 |  | 1
| 1 | 3 | 0.7 | 0.61 | 0.61 | 0.49 | 0.54 | 0.43972 | 0.5 | 0.41 | 0.65 | 0.4 | 0.67567 | 0.8 | 0.54 |  | 16
| 1 | 4 | 0.72 | 0.56 | 0.56 | 0.48 | 0.53 | 0.45224 | 0.4 | 0.42 | 0.638 | 0.4 | 0.62642 | 0.5 | 0.52 |  | 0.52948
| 1 | 5 | 0.72 | 0.54 | 0.6 | 0.46 | 0.54 | 0.39683 | 0.4 | 0.43 | 0.637 | 0.43 | 0.6078 | 0.5 | 0.54 |  | 91
| 1 | 6 | 0.69 | 0.58 | 0.57 | 0.48 | 0.56 | 0.40016 | 0.4 | 0.38 | 0.637 | 0.43 | 0.57718 | 0.5 | 0.54
| Mean |  | 0.72167 | 0.56833 | 0.58333 | 0.48333 | 0.54667 | 0.42935 | 0.45 | 0.4 | 0.64183 | 0.41333 | 0.61867 | 0.55 | 0.52667
| Median |  | 0.72 | 0.565 | 0.585 | 0.485 | 0.54 | 0.43624 | 0.45 | 0.405 | 0.638 | 0.41 | 0.61684 | 0.5 | 0.53
| Std Dev. |  | 0.025626 | 0.024833 | 0.022509 | 0.013663 | 0.015055 | 0.025221 | 0.054772 | 0.026077 | 0.0067355 | 0.013663 | 0.033423 | 0.12247 | 0.01633
| Rel.Std.Dev. |  | 3.551% | 4.369% | 3.859% | 2.827% | 2.754% | 5.874% | 12.17% | 6.519% | 1.049% | 3.305% | 5.402% | 22.27% | 3.101%
| PDM3 |  | 36.3% | 7.339% | 10.17% | -8.715% | 3.247% | -18.91% | -15.01% | -24.45% | 21.22% | -21.94% | 16.85% | 3.876% | -0.5306%
| Z-Score (Absolute) |  | 1.72 | 0.37 | 0.5 | 0.38 | 0.18 | 0.86 | 0.67 | 1.11 | 1.01 | 1 | 0.81 | 0.23 | 0
| Analytical results for Mg in OREAS 462 (Certified Value 1.01 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
2
| 1 | 1 | 0.95 | 0.96 | 0.75 | 0.78 | 0.71 | 1.1546 | 1.02 | 1.08 | 1.08 | 1.1381 | 1.0533 | 0.97 | 1.0855 | 1.1 | 1
| 1 | 2 | 0.96 | 0.95 | 0.76 | 0.81 | 0.72 | 1.1308 | 1.02 | 1.07 | 1.09 | 1.1719 | 1.0527 | 0.98 | 1.075 | 1.1 | 2
| 1 | 3 | 0.87 | 0.99 | 0.76 | 0.82 | 0.72 | 1.136 | 1.01 | 1.07 | 1.08 | 1.119 | 1.0687 | 0.99 | 1.07 | 1.11 | 16
| 1 | 4 | 0.9 | 0.94 | 0.77 | 0.81 | 0.71 | 1.1503 | 1.02 | 1.06 | 1.09 | 1.144 | 1.0672 | 0.98 | 1.0775 | 1.11 | 1.007
| 1 | 5 | 0.92 | 0.94 | 0.76 | 0.79 | 0.8 | 1.1516 | 1.02 | 1.05 | 1.07 | 1.118 | 1.0541 | 0.98 | 1.0735 | 1.14 | 92
| 1 | 6 | 0.87 | 0.96 | 0.77 | 0.79 | 0.73 | 1.1242 | 1.01 | 1.05 | 1.07 | 1.1136 | 1.0538 | 0.95 | 1.071 | 1.11
| Mean |  | 0.91167 | 0.95667 | 0.76167 | 0.8 | 0.73167 | 1.1413 | 1.0167 | 1.0633 | 1.08 | 1.1341 | 1.0583 | 0.975 | 1.0754 | 1.1117
| Median |  | 0.91 | 0.955 | 0.76 | 0.8 | 0.72 | 1.1432 | 1.02 | 1.065 | 1.08 | 1.1285 | 1.054 | 0.98 | 1.0742 | 1.11
| Std Dev. |  | 0.038687 | 0.018619 | 0.0075277 | 0.015492 | 0.034303 | 0.012607 | 0.005164 | 0.012111 | 0.0089443 | 0.022137 | 0.0075158 | 0.013784 | 0.005634 | 0.01472
| Rel.Std.Dev. |  | 4.244% | 1.946% | 0.9883% | 1.936% | 4.688% | 1.105% | 0.5079% | 1.139% | 0.8282% | 1.952% | 0.7102% | 1.414% | 0.5239% | 1.324%
| PDM3 |  | -9.465% | -4.997% | -24.36% | -20.55% | -27.34% | 13.33% | 0.9618% | 5.596% | 7.251% | 12.62% | 5.098% | -3.176% | 6.796% | 10.4%
| Z-Score (Absolute) |  | 1.09 | 0.7 | 2.4 | 2.07 | 2.78 | 0.9 | 0.18 | 0.22 | 0.37 | 0.84 | 0.18 | 0.5 | 0.33 | 0.65
| Analytical results for Mn in OREAS 462 (Certified Value 840 ppm) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 767 | 810 | 541 | 534 | 541 | 854.1404 | 898.3 | 820 | 816 | 911 | 1288 | 860 | 944 |  | 1
| 1 | 2 | 745 | 787 | 538 | 574 | 520 | 845.835 | 883.1 | 787 | 807 | 920 | 1256 | 850 | 938 |  | 3
| 1 | 3 | 735 | 817 | 545 | 586 | 520 | 827.5283 | 880 | 809 | 802 | 889 | 1231 | 845 | 934 |  | 16
| 1 | 4 | 711 | 790 | 545 | 580 | 557 | 845.6005 | 875.4 | 808 | 809 | 923 | 1244 | 850 | 956 |  | 839.9573
| 1 | 5 | 734 | 769 | 538 | 572 | 534 | 828.8735 | 872.1 | 795 | 794 | 896 | 1265 | 845 | 978 |  | 93
| 1 | 6 | 732 | 772 | 538 | 553 | 574 | 829.214 | 864.6 | 793 | 801 | 894 | 1252 | 850 | 960
| Mean |  | 737.3333 | 790.8333 | 540.8333 | 566.5 | 541 | 838.532 | 878.9167 | 802 | 804.8333 | 905.5 | 1256 | 850 | 951.6667
| Median |  | 734.5 | 788.5 | 539.5 | 573 | 537.5 | 837.4073 | 877.7 | 801.5 | 804.5 | 903.5 | 1254 | 850 | 950
| Std Dev. |  | 18.3157 | 19.4876 | 3.4303 | 19.4294 | 21.3354 | 11.3852 | 11.4733 | 12.3612 | 7.5741 | 14.4326 | 19.4422 | 5.4772 | 16.3666
| Rel.Std.Dev. |  | 2.484% | 2.464% | 0.6343% | 3.43% | 3.944% | 1.358% | 1.305% | 1.541% | 0.9411% | 1.594% | 1.548% | 0.6444% | 1.72%
| PDM3 |  | -12.22% | -5.848% | -35.61% | -32.56% | -35.59% | -0.1697% | 4.638% | -4.519% | -4.182% | 7.803% | 49.53% | 1.196% | 13.3%
| Z-Score (Absolute) |  | 0.61 | 0.13 | 2.4 | 2.17 | 2.4 | 0.31 | 0.67 | 0.03 | 0 | 0.92 | 4.11 | 0.41 | 1.34
| Analytical results for Mo in OREAS 462 (Certified Value 53 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
1
| 1 | 1 | 52.1 | 53.1 | 61.9 | 55.5 | 55.1 | 50.3704 | 51.5 | 53 | 54.4 | 54.4 | 50.8601 | 46 |  |  | 1
| 1 | 2 | 52.3 | 54.2 | 60.2 | 55.6 | 58.9 | 50.3981 | 52.1 | 54 | 53.5 | 54.1 | 51.8125 | 47.5 |  |  | 0
| 1 | 3 | 52.4 | 58.7 | 63 | 55.4 | 56.4 | 48.8612 | 51.5 | 56 | 53.4 | 54.7 | 50.4423 | 47 |  |  | 16
| 1 | 4 | 51 | 53.5 | 60.8 | 55.1 | 55.5 | 50.2379 | 51.4 | 57 | 54.2 | 53.4 | 50.8098 | 47.5 |  |  | 53.4403
| 1 | 5 | 51.5 | 55.2 | 63.2 | 54 | 52.8 | 49.1176 | 51.3 | 55 | 53.5 | 54.5 | 51.9794 | 47.5 |  |  | 94
| 1 | 6 | 51 | 54.7 | 59.2 | 56.3 | 56.2 | 48.6957 | 51.8 | 54 | 53.7 | 55.1 | 51.9743 | 47.5
| Mean |  | 51.7167 | 54.9 | 61.3833 | 55.3167 | 55.8167 | 49.6135 | 51.6 | 54.8333 | 53.7833 | 54.3667 | 51.3131 | 47.1667
| Median |  | 51.8 | 54.45 | 61.35 | 55.45 | 55.85 | 49.6778 | 51.5 | 54.5 | 53.6 | 54.45 | 51.3363 | 47.5
| Std Dev. |  | 0.63692 | 2.013 | 1.593 | 0.75741 | 1.9854 | 0.80405 | 0.29665 | 1.472 | 0.41673 | 0.5785 | 0.68518 | 0.60553
| Rel.Std.Dev. |  | 1.232% | 3.667% | 2.595% | 1.369% | 3.557% | 1.621% | 0.5749% | 2.684% | 0.7748% | 1.064% | 1.335% | 1.284%
| PDM3 |  | -3.225% | 2.732% | 14.86% | 3.511% | 4.447% | -7.161% | -3.444% | 2.607% | 0.642% | 1.734% | -3.981% | -11.74%
| Z-Score (Absolute) |  | 0.74 | 0.06 | 2.44 | 0.45 | 0.61 | 1.43 | 0.78 | 0.29 | 0.06 | 0.13 | 0.87 | 2.16
| Analytical results for Na in OREAS 462 (Certified Value 0.178 wt.%) | [15 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.18 | 0.194 | 0.1872 | 0.18416 | 0.19 | 0.17 | 0.2 |  | 1
| 1 | 2 | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.18 | 0.195 | 0.1925 | 0.18536 | 0.18 | 0.17 | 0.19 |  | 4
| 1 | 3 | 0.16 | 0.16 | 0.17 | 0.19 | 0.18 | 0.18 | 0.17 | 0.194 | 0.1815 | 0.18831 | 0.18 | 0.17 | 0.19 |  | 16
| 1 | 4 | 0.16 | 0.16 | 0.17 | 0.19 | 0.16 | 0.18 | 0.17 | 0.196 | 0.1841 | 0.19614 | 0.18 | 0.17 | 0.2 |  | 0.17785
| 1 | 5 | 0.16 | 0.16 | 0.17 | 0.18 | 0.2 | 0.18 | 0.17 | 0.191 | 0.1821 | 0.18727 | 0.18 | 0.17 | 0.2 |  | 95
| 1 | 6 | 0.15 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.17 | 0.192 | 0.1786 | 0.18297 | 0.18 | 0.17 | 0.19
| Mean |  | 0.15833 | 0.16 | 0.17 | 0.18833 | 0.175 | 0.18 | 0.17333 | 0.19367 | 0.18433 | 0.18737 | 0.18167 | 0.17 | 0.195
| Median |  | 0.16 | 0.16 | 0.17 | 0.19 | 0.17 | 0.18 | 0.17 | 0.194 | 0.1831 | 0.18631 | 0.18 | 0.17 | 0.195
| Std Dev. |  | 0.0040825 | 0 | 0 | 0.0040825 | 0.013784 | 3.0405e-17 | 0.005164 | 0.0018619 | 0.0049164 | 0.0047225 | 0.0040825 | 3.0405e-17 | 0.0054772
| Rel.Std.Dev. |  | 2.578% | 0% | 0% | 2.168% | 7.877% | 1.689e-14% | 2.979% | 0.9614% | 2.667% | 2.52% | 2.247% | 1.789e-14% | 2.809%
| PDM3 |  | -10.97% | -10.04% | -4.413% | 5.895% | -1.602% | 1.21% | -2.539% | 8.894% | 3.646% | 5.351% | 2.147% | -4.413% | 9.644%
| Z-Score (Absolute) |  | 1.46 | 1.35 | 0.67 | 0.56 | 0.67 | 0 | 0.45 | 0.92 | 0.29 | 0.5 | 0.11 | 0.67 | 1.01
| Analytical results for Nb in OREAS 462 (Indicative Value 1231 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 1239.1475 | 1253.5 | 1002 | > 500 | 1410 |  |  |  |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 1271.2706 | 1257.1 | 997 | > 500 | 1460 |  |  |  |  | 5
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 1244.0739 | 1222.3 | 1004 | > 500 | 1440 |  |  |  |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 1266.313 | 1226.8 | 953 | > 500 | 1480 |  |  |  |  | 1231.0306
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 1234.1294 | 1224.3 | 994 | > 500 | 1450 |  |  |  |  | 27
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 1251 | 1232.8 | 982 | > 500 | 1450
| Mean |  | 07: 1250.9891; 08: 1236.1333; 09: 988.6667; 22: 1448.3333
| Median |  | 07: 1247.5369; 08: 1229.8; 09: 995.5; 22: 1450
| Std Dev. |  | 07: 14.9563; 08: 15.3024; 09: 19.1172; 22: 23.1661
| Rel.Std.Dev. |  | 07: 1.196%; 08: 1.238%; 09: 1.934%; 22: 1.599%
| PDM3 |  | 07: 1.621%; 08: 0.4145%; 09: -19.69%; 22: 17.65%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.05 | 0.05 | 1.62 | NA | 1.3
| Analytical results for Nd in OREAS 462 (Certified Value 2595 ppm) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 2620 | 2650 | 2640 | 2690 | 2420 | 2838.8944 | 2534 | 2557 | 2710 | 2417.59 | 2470.2447 | 2672 | 2460 |  | 1
| 1 | 2 | 2690 | 2680 | 2500 | 2660 | 2590 | 2869.9753 | 2502.6 | 2577 | 2710 | 2485.15 | 2479.605 | 2670 | 2600 |  | 0
| 1 | 3 | 2440 | 3000 | 2750 | 2700 | 2480 | 2873.9858 | 2485.5 | 2554 | 2740 | 2455.36 | 2471.7329 | 2650 | 2620 |  | 16
| 1 | 4 | 2580 | 2720 | 2690 | 2660 | 2390 | 2755.2981 | 2474 | 2540 | 2740 | 2416.43 | 2424.0226 | 2672.5 | 2580 |  | 2595.2838
| 1 | 5 | 2490 | 2620 | 2680 | 2640 | 2430 | 2791.6196 | 2506.7 | 2588 | 2710 | 2443.19 | 2503.0454 | 2639.5 | 2550 |  | 96
| 1 | 6 | 2400 | 2760 | 2540 | 2740 | 2440 | 2827.4021 | 2467.3 | 2551 | 2740 | 2426.66 | 2458.3275 | 2698.5 | 2580
| Mean |  | 2536.6667 | 2738.3333 | 2633.3333 | 2681.6667 | 2458.3333 | 2826.1959 | 2495.0167 | 2561.1667 | 2725 | 2440.73 | 2467.8297 | 2667.0833 | 2565
| Median |  | 2535 | 2700 | 2660 | 2675 | 2435 | 2833.1483 | 2494.05 | 2555.5 | 2725 | 2434.925 | 2470.9888 | 2671 | 2580
| Std Dev. |  | 111.8332 | 137.4651 | 95.4289 | 36.0093 | 70.8284 | 46.0114 | 24.5528 | 17.8372 | 16.4317 | 26.5222 | 26.1122 | 20.5022 | 56.4801
| Rel.Std.Dev. |  | 4.409% | 5.02% | 3.624% | 1.343% | 2.881% | 1.628% | 0.9841% | 0.6964% | 0.603% | 1.087% | 1.058% | 0.7687% | 2.202%
| PDM3 |  | -2.259% | 5.512% | 1.466% | 3.328% | -5.277% | 8.897% | -3.863% | -1.315% | 4.998% | -5.955% | -4.911% | 2.767% | -1.167%
| Z-Score (Absolute) |  | 0.35 | 0.7 | 0.33 | 0.67 | 1.09 | 1.69 | 0.64 | 0.18 | 0.98 | 1.02 | 0.83 | 0.57 | 0
| Analytical results for Ni in OREAS 462 (Certified Value 74 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
1
| 1 | 1 | 74 | 72.7 | 74.7 | 74.1 | 68.8 | 68.9525 | 78 | 77.2 | 77 | 75 | 90 |  |  |  | 1
| 1 | 2 | 75 | 72.2 | 72.6 | 72.8 | 74 | 70.5688 | 81 | 74 | 79 | 70 | 88 |  |  |  | 0
| 1 | 3 | 75 | 78.1 | 78.6 | 73.1 | 71.1 | 69.4166 | 87 | 75.4 | 77 | 75 | 92 |  |  |  | 16
| 1 | 4 | 73 | 70.8 | 76.7 | 72.9 | 70.8 | 71.0699 | 82 | 73.1 | 75 | 75 | 98 |  |  |  | 73.8248
| 1 | 5 | 75 | 73.1 | 77.1 | 71.8 | 69.7 | 69.6211 | 73 | 75.1 | 75 | 70 | 96 |  |  |  | 97
| 1 | 6 | 72 | 71.8 | 73 | 73.5 | 70 | 70.2364 | 76 | 75.8 | 81 | 70 | 92
| Mean |  | 74 | 73.1167 | 75.45 | 73.0333 | 70.7333 | 69.9776 | 79.5 | 75.1 | 77.3333 | 72.5 | 92.6667
| Median |  | 74.5 | 72.45 | 75.7 | 73 | 70.4 | 69.9288 | 79.5 | 75.25 | 77 | 72.5 | 92
| Std Dev. |  | 1.2649 | 2.567 | 2.404 | 0.76855 | 1.7974 | 0.78769 | 4.9295 | 1.4283 | 2.3381 | 2.7386 | 3.7238
| Rel.Std.Dev. |  | 1.709% | 3.511% | 3.186% | 1.052% | 2.541% | 1.126% | 6.201% | 1.902% | 3.023% | 3.777% | 4.018%
| PDM3 |  | 0.2374% | -0.9591% | 2.201% | -1.072% | -4.188% | -5.211% | 7.687% | 1.727% | 4.753% | -1.794% | 25.52%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.52 | 0.35 | 1.17 | 1.44 | 1.43 | 0.39 | 1.2 | 0.54 | 6.7
| Analytical results for P in OREAS 462 (Certified Value 0.454 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.44 | 0.473 | 0.365 | 0.391 | 0.357 | 0.4636 | 0.51 | 0.479 | 0.4832 | 0.52894 | 0.4396 | 0.488 | 0.48 |  | 1
| 1 | 2 | 0.44 | 0.468 | 0.368 | 0.405 | 0.362 | 0.4574 | 0.49 | 0.486 | 0.4963 | 0.52717 | 0.4311 | 0.484 | 0.475 |  | 6
| 1 | 3 | 0.431 | 0.484 | 0.371 | 0.406 | 0.364 | 0.4592 | 0.51 | 0.482 | 0.4868 | 0.52193 | 0.4407 | 0.486 | 0.51 |  | 16
| 1 | 4 | 0.429 | 0.465 | 0.371 | 0.399 | 0.357 | 0.4667 | 0.51 | 0.488 | 0.491 | 0.52031 | 0.431 | 0.483 | 0.48 |  | 0.45372
| 1 | 5 | 0.438 | 0.464 | 0.369 | 0.393 | 0.398 | 0.4675 | 0.51 | 0.474 | 0.4786 | 0.51945 | 0.4346 | 0.479 | 0.515 |  | 98
| 1 | 6 | 0.425 | 0.471 | 0.373 | 0.396 | 0.368 | 0.4617 | 0.5 | 0.479 | 0.4828 | 0.5221 | 0.4228 | 0.473 | 0.48
| Mean |  | 0.43383 | 0.47083 | 0.3695 | 0.39833 | 0.36767 | 0.46268 | 0.505 | 0.48133 | 0.48645 | 0.52332 | 0.4333 | 0.48217 | 0.49
| Median |  | 0.4345 | 0.4695 | 0.37 | 0.3975 | 0.363 | 0.46265 | 0.51 | 0.4805 | 0.485 | 0.52201 | 0.43285 | 0.4835 | 0.48
| Std Dev. |  | 0.0063692 | 0.0073052 | 0.0028107 | 0.006186 | 0.015449 | 0.0040296 | 0.0083666 | 0.0051251 | 0.0063774 | 0.0038422 | 0.0065806 | 0.0054191 | 0.017607
| Rel.Std.Dev. |  | 1.468% | 1.552% | 0.7607% | 1.553% | 4.202% | 0.8709% | 1.657% | 1.065% | 1.311% | 0.7342% | 1.519% | 1.124% | 3.593%
| PDM3 |  | -4.383% | 3.772% | -18.56% | -12.21% | -18.97% | 1.976% | 11.3% | 6.086% | 7.214% | 15.34% | -4.5% | 6.27% | 7.996%
| Z-Score (Absolute) |  | 0.73 | 0 | 2 | 1.43 | 2.16 | 0.16 | 0.67 | 0.21 | 0.31 | 1.04 | 0.74 | 0.22 | 0.38
| Analytical results for Pb in OREAS 462 (Certified Value 119 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 111 | 116 | 132 | 121.5 | 116.5 | 119.055 | 117 | 121 | 95.1 | 175.525 | 160 | 113 |  |  | 1
| 1 | 2 | 112 | 118.5 | 125 | 120.5 | 123.5 | 120.7467 | 118 | 121 | 98.7 | 175.738 | 150 | 112 |  |  | 0
| 1 | 3 | 110 | 122.5 | 130 | 122 | 121 | 119.951 | 114 | 119 | 138.6 | 175.776 | 160 | 114 |  |  | 16
| 1 | 4 | 111 | 113.5 | 122 | 119.5 | 117 | 121.6484 | 115 | 120 | 96.4 | 175.853 | 150 | 108 |  |  | 118.5188
| 1 | 5 | 111 | 113.5 | 129.5 | 117 | 128 | 120.0834 | 117 | 119 | 95.7 | 174.521 | 150 | 110 |  |  | 99
| 1 | 6 | 107 | 120.5 | 128.5 | 122 | 118.5 | 121.0328 | 118 | 120 | 97 | 175.704 | 150 | 117
| Mean |  | 110.3333 | 117.4167 | 127.8333 | 120.4167 | 120.75 | 120.4195 | 116.5 | 120 | 103.5833 | 175.5195 | 153.3333 | 112.3333
| Median |  | 111 | 117.25 | 129 | 121 | 119.75 | 120.415 | 117 | 120 | 96.7 | 175.721 | 150 | 112.5
| Std Dev. |  | 1.7512 | 3.7204 | 3.6697 | 1.9343 | 4.413 | 0.91521 | 1.6432 | 0.89443 | 17.1992 | 0.50116 | 5.164 | 3.1411
| Rel.Std.Dev. |  | 1.587% | 3.169% | 2.871% | 1.606% | 3.655% | 0.76% | 1.41% | 0.7454% | 16.6% | 0.2855% | 3.368% | 2.796%
| PDM3 |  | -6.907% | -0.93% | 7.859% | 1.601% | 1.883% | 1.604% | -1.703% | 1.25% | -12.6% | 48.09% | 29.37% | -5.219%
| Z-Score (Absolute) |  | 1.1 | 0.33 | 0.91 | 0.02 | 0.06 | 0.03 | 0.44 | 0.02 | 2.81 | 6.58 | 3.94 | 0.94
| Analytical results for Pr in OREAS 462 (Certified Value 736 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 796 | 700 | 771 | 746 | 676 | 736.9686 | 725.4 | 724.3 | 734 | 709.85 | 768.4469 | 748.1 | 694 |  | 1
| 1 | 2 | 812 | 713 | 725 | 737 | 730 | 787.9018 | 714.1 | 728.19 | 729 | 735.52 | 765.1519 | 747.6 | 715 |  | 0
| 1 | 3 | 730 | 784 | 809 | 753 | 698 | 793.0289 | 716.5 | 719.42 | 739 | 725.77 | 767.6484 | 743.1 | 724 |  | 16
| 1 | 4 | 770 | 724 | 776 | 734 | 670 | 772.8533 | 704.5 | 717.51 | 739 | 713 | 768.0107 | 749.4 | 718 |  | 735.9837
| 1 | 5 | 753 | 697 | 782 | 731 | 683 | 737.9657 | 712.7 | 732.91 | 739 | 725.44 | 765.6294 | 740.5 | 715 |  | 100
| 1 | 6 | 716 | 720 | 737 | 759 | 689 | 789.6156 | 706.7 | 715.1 | 743 | 734.53 | 767.3691 | 752.2 | 707
| Mean |  | 762.8333 | 723 | 766.6667 | 743.3333 | 691 | 769.7223 | 713.3167 | 722.905 | 737.1667 | 724.0183 | 767.0427 | 746.8167 | 712.1667
| Median |  | 761.5 | 716.5 | 773.5 | 741.5 | 686 | 780.3776 | 713.4 | 721.86 | 739 | 725.605 | 767.5088 | 747.85 | 715
| Std Dev. |  | 37.2474 | 31.7364 | 30.8199 | 11.1833 | 21.4663 | 25.925 | 7.4652 | 6.8139 | 4.916 | 10.6772 | 1.3383 | 4.2808 | 10.4579
| Rel.Std.Dev. |  | 4.883% | 4.39% | 4.02% | 1.504% | 3.107% | 3.368% | 1.047% | 0.9426% | 0.6669% | 1.475% | 0.1745% | 0.5732% | 1.468%
| PDM3 |  | 3.648% | -1.764% | 4.169% | 0.9986% | -6.112% | 4.584% | -3.08% | -1.777% | 0.1607% | -1.626% | 4.22% | 1.472% | -3.236%
| Z-Score (Absolute) |  | 0.69 | 0.71 | 0.8 | 0.17 | 1.25 | 0.88 | 0.64 | 0.38 | 0 | 0.35 | 0.81 | 0.26 | 0.67
| Analytical results for Rb in OREAS 462 (Certified Value 6.45 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 6.8 | 6.6 | 7 | 6.4 | 5.9 | 6.1826 | 6.3 | 6.7 | 5.34 | 6.7 | 7.2882 | 6 | 6.8 |  | 1
| 1 | 2 | 7.3 | 6.7 | 6.9 | 6.2 | 6.4 | 6.17 | 6.1 | 7.1 | 5.27 | 6.68 | 7.3447 | 6 | 6.8 |  | 0
| 1 | 3 | 3.3 | 7.3 | 7.1 | 6.1 | 6 | 6.0464 | 6.1 | 6.9 | 5.3 | 6.66 | 7.8379 | 6 | 6.8 |  | 16
| 1 | 4 | 4.9 | 6.7 | 7 | 6.2 | 5.9 | 6.1962 | 6.3 | 7 | 5.34 | 6.39 | 7.1816 | 6.5 | 6.8 |  | 6.4463
| 1 | 5 | 3.9 | 6.7 | 7.3 | 6.1 | 5.8 | 5.9884 | 6.2 | 7.2 | 5.26 | 6.43 | 7.6916 | 6 | 6.4 |  | 101
| 1 | 6 | 3.6 | 6.8 | 7 | 6.3 | 5.9 | 6.2513 | 6.3 | 7 | 5.34 | 6.59 | 6.9337 | 6 | 6.8
| Mean |  | 4.9667 | 6.8 | 7.05 | 6.2167 | 5.9833 | 6.1391 | 6.2167 | 6.9833 | 5.3083 | 6.575 | 7.3796 | 6.0833 | 6.7333
| Median |  | 4.4 | 6.7 | 7 | 6.2 | 5.9 | 6.1763 | 6.25 | 7 | 5.32 | 6.625 | 7.3164 | 6 | 6.8
| Std Dev. |  | 1.7084 | 0.25298 | 0.13784 | 0.1169 | 0.2137 | 0.099999 | 0.098319 | 0.17224 | 0.037103 | 0.13368 | 0.33316 | 0.20412 | 0.1633
| Rel.Std.Dev. |  | 34.4% | 3.72% | 1.955% | 1.881% | 3.572% | 1.629% | 1.582% | 2.466% | 0.699% | 2.033% | 4.515% | 3.355% | 2.425%
| PDM3 |  | -22.95% | 5.487% | 9.365% | -3.562% | -7.182% | -4.765% | -3.562% | 8.331% | -17.65% | 1.996% | 14.48% | -5.631% | 4.452%
| Z-Score (Absolute) |  | 0 | 0.9 | 1.38 | 0.22 | 0.67 | 0.37 | 0.22 | 1.25 | 1.97 | 0.47 | 2.02 | 0.48 | 0.77
| Analytical results for Re in OREAS 462 (Certified Value < 0.01 ppm) | [15 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.014 | < 0.002 | 0.005 | 0.005 | 0.005 | < 0.002 | < 0.02 | 0.004 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.012 | 0.005 | 0.006 | 0.003 | 0.005 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 18
| 1 | 3 | 0.012 | 0.004 | 0.006 | 0.004 | 0.006 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.012 | < 0.002 | 0.004 | 0.005 | 0.005 | < 0.002 | < 0.02 | 0.004 | < 0.002 | < 0.05 | < 0.1 |  |  |  | < 0.01
| 1 | 5 | 0.011 | 0.004 | 0.005 | 0.004 | 0.006 | < 0.002 | < 0.02 | 0.004 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 102
| 1 | 6 | 0.012 | < 0.002 | 0.004 | 0.006 | 0.005 | < 0.002 | < 0.02 | 0.006 | < 0.002 | < 0.05 | < 0.1
| Mean |  | 01: 0.012167; 02: 0.0043333; 03: 0.005; 05: 0.0045; 06: 0.0053333; 09: 0.004
| Median |  | 01: 0.012; 02: 0.004; 03: 0.005; 05: 0.0045; 06: 0.005; 09: 0.004
| Std Dev. |  | 01: 0.00098319; 02: 0.00057735; 03: 0.00089443; 05: 0.0010488; 06: 0.0005164; 09: 0.0010954
| Rel.Std.Dev. |  | 01: 8.081%; 02: 13.32%; 03: 17.89%; 05: 23.31%; 06: 9.682%; 09: 27.39%
PDM3
| Z-Score (Absolute) |  | 7.13 | 2.11 | 0.25 | 0.25 | 0.59 | NA | NA | 0.76 | NA | NA | NA
| Analytical results for S in OREAS 462 (Certified Value 548 ppm) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 500 | 600 | 500 | 600 | 500 | 500 | 500 | 550 | 625 | 400 | 600 | 600 |  |  | 1
| 1 | 2 | 500 | 600 | 500 | 600 | 600 | 500 | 500 | 550 | 631 | 400 | 600 | 650 |  |  | 9
| 1 | 3 | 500 | 600 | 500 | 600 | 600 | 500 | 500 | 540 | 579 | 400 | 600 | 600 |  |  | 16
| 1 | 4 | 500 | 600 | 500 | 600 | 500 | 500 | 500 | 550 | 708 | 400 | 600 | 600 |  |  | 547.9444
| 1 | 5 | 500 | 600 | 500 | 600 | 600 | 500 | 500 | 540 | 671 | 400 | 600 | 650 |  |  | 103
| 1 | 6 | 500 | 600 | 500 | 600 | 600 | 500 | 500 | 540 | 618 | 400 | 600 | 650
| Mean |  | 500 | 600 | 500 | 600 | 566.6667 | 500 | 500 | 545 | 638.6667 | 400 | 600 | 625
| Median |  | 500 | 600 | 500 | 600 | 600 | 500 | 500 | 545 | 628 | 400 | 600 | 625
| Std Dev. |  | 0 | 0 | 0 | 0 | 51.6398 | 0 | 0 | 5.4772 | 44.9118 | 0 | 0 | 27.3861
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 9.113% | 0% | 0% | 1.005% | 7.032% | 0% | 0% | 4.382%
| PDM3 |  | -8.75% | 9.5% | -8.75% | 9.5% | 3.417% | -8.75% | -8.75% | -0.5374% | 16.56% | -27% | 9.5% | 14.06%
| Z-Score (Absolute) |  | 0.67 | 0.53 | 0.67 | 0.53 | 0.13 | 0.67 | 0.67 | 0.13 | 1 | 1.88 | 0.53 | 0.84
| Analytical results for Sb in OREAS 462 (Certified Value 2.17 ppm) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.17 | 2.32 | 2.63 | 2.29 | 2.25 | 1.9149 | 2.2 | 1.6 | 2.2 | 1.7 | 1.6 |  |  |  | 1
| 1 | 2 | 2.28 | 2.44 | 2.49 | 2.36 | 2.41 | 1.9739 | 2.2 | 1.9 | 2.15 | 1.59 | 2 |  |  |  | 0
| 1 | 3 | 2.13 | 2.64 | 2.66 | 2.39 | 2.31 | 1.9711 | 2.3 | 2.1 | 2.14 | 1.78 | 1.8 |  |  |  | 16
| 1 | 4 | 2.21 | 2.45 | 2.64 | 2.32 | 2.23 | 1.895 | 2.2 | 2.3 | 2.21 | 1.81 | 1.8 |  |  |  | 2.172
| 1 | 5 | 2.05 | 2.39 | 2.71 | 2.34 | 2.25 | 1.9478 | 2.3 | 2.1 | 2.12 | 2.03 | 2 |  |  |  | 104
| 1 | 6 | 2.12 | 2.38 | 2.53 | 2.39 | 2.3 | 1.9958 | 2.3 | 2.1 | 2.16 | 1.79 | 2.6
| Mean |  | 2.16 | 2.4367 | 2.61 | 2.3483 | 2.2917 | 1.9497 | 2.25 | 2.0167 | 2.1633 | 1.7833 | 1.9667
| Median |  | 2.15 | 2.415 | 2.635 | 2.35 | 2.275 | 1.9594 | 2.25 | 2.1 | 2.155 | 1.785 | 1.9
| Std Dev. |  | 0.079498 | 0.11003 | 0.083187 | 0.039707 | 0.065853 | 0.038408 | 0.054772 | 0.24014 | 0.035024 | 0.14556 | 0.34448
| Rel.Std.Dev. |  | 3.68% | 4.516% | 3.187% | 1.691% | 2.874% | 1.97% | 2.434% | 11.91% | 1.619% | 8.162% | 17.52%
| PDM3 |  | -0.5542% | 12.18% | 20.16% | 8.117% | 5.508% | -10.23% | 3.589% | -7.153% | -0.4008% | -17.9% | -9.455%
| Z-Score (Absolute) |  | 0.01 | 0.85 | 1.63 | 0.67 | 0.47 | 0.78 | 0.32 | 0.23 | 0 | 1.39 | 1.18
| Analytical results for Sc in OREAS 462 (Certified Value 51 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 0
1
| 1 | 1 | 51.3 | 55.2 | 54.9 | 53.3 | 50.4 | 50.9321 | 56 | 49 | 44.2 | 54.6519 | 57 | 51 |  |  | 1
| 1 | 2 | 51.3 | 55.9 | 51.6 | 52.7 | 54.3 | 50.069 | 57 | 47 | 44.7 | 54.972 | 56 | 50 |  |  | 19
| 1 | 3 | 43.5 | 60.7 | 56.6 | 52.7 | 52.3 | 47.5553 | 53 | 49 | 45.9 | 54.3046 | 55 | 50 |  |  | 16
| 1 | 4 | 46.1 | 56.1 | 54 | 52.1 | 51 | 48.8575 | 53 | 47 | 45 | 54.8559 | 56 | 50 |  |  | 51.4401
| 1 | 5 | 45.3 | 53.6 | 55.5 | 51.6 | 51 | 46.9632 | 51 | 48 | 46 | 54.793 | 55 | 52 |  |  | 105
| 1 | 6 | 43.2 | 56.9 | 52.7 | 53.4 | 51.2 | 49.2741 | 51 | 48 | 46 | 54.538 | 56 | 51
| Mean |  | 46.7833 | 56.4 | 54.2167 | 52.6333 | 51.7 | 48.9419 | 53.5 | 48 | 45.3 | 54.6859 | 55.8333 | 50.6667
| Median |  | 45.7 | 56 | 54.45 | 52.7 | 51.1 | 49.0658 | 53 | 48 | 45.45 | 54.7225 | 56 | 50.5
| Std Dev. |  | 3.663 | 2.3816 | 1.8433 | 0.69186 | 1.417 | 1.4958 | 2.51 | 0.89443 | 0.7746 | 0.24109 | 0.75277 | 0.8165
| Rel.Std.Dev. |  | 7.83% | 4.223% | 3.4% | 1.314% | 2.741% | 3.056% | 4.692% | 1.863% | 1.71% | 0.4409% | 1.348% | 1.612%
| PDM3 |  | -9.053% | 9.642% | 5.398% | 2.32% | 0.5053% | -4.857% | 4.004% | -6.688% | -11.94% | 6.31% | 8.54% | -1.504%
| Z-Score (Absolute) |  | 1.2 | 0.85 | 0.54 | 0.17 | 0.17 | 0.7 | 0.37 | 0.92 | 1.55 | 0.65 | 0.92 | 0.29
| Analytical results for Se in OREAS 462 (Indicative Value 8.24 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 5 | < 50 | 7 | 8 | 12 | 2 | 12 | 10 | < 0.5 | < 0.5 | 10 |  |  |  | 1
| 1 | 2 | 5 | < 50 | 7 | 8 | 13 | 1.5937 | 12 | 10 | < 0.5 | < 0.5 | 10 |  |  |  | 15
| 1 | 3 | 5 | < 50 | 6 | 9 | 12 | 2 | 13 | 9 | < 0.5 | < 0.5 | 10 |  |  |  | 16
| 1 | 4 | 5 | < 50 | 7 | 9 | 12 | 1.4224 | 12 | 10 | < 0.5 | < 0.5 | 10 |  |  |  | 8.2354
| 1 | 5 | 5 | < 50 | 8 | 8 | 12 | 1.5 | 12 | 9 | < 0.5 | < 0.5 | 10 |  |  |  | 27
| 1 | 6 | 5 | < 50 | 7 | 8 | 12 | 1.7848 | 12 | 9 | < 0.5 | < 0.5 | 10
| Mean |  | 5 |  | 7 | 8.3333 | 12.1667 | 1.7168 | 12.1667 | 9.5 |  |  | 10
| Median |  | 5 |  | 7 | 8 | 12 | 1.6893 | 12 | 9.5 |  |  | 10
| Std Dev. |  | 0 |  | 0.63246 | 0.5164 | 0.40825 | 0.25054 | 0.40825 | 0.54772 |  |  | 0
| Rel.Std.Dev. |  | 0% |  | 9.035% | 6.197% | 3.355% | 14.59% | 3.355% | 5.766% |  |  | 0%
| PDM3 |  | -39.29% |  | -15% | 1.189% | 47.74% | -79.15% | 47.74% | 15.36% |  |  | 21.43%
| Z-Score (Absolute) |  | 1.02 | NA | 0.5 | 0.15 | 0.85 | 1.88 | 0.85 | 0.15 | NA | NA | 0.28
| Analytical results for Sm in OREAS 462 (Certified Value 362 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 349 | 366 | 367 | 357 | 332 | 393.3751 | 373.4 | 350.2 | 352 | 341.17 | 384.4328 | 361.3 | 366 |  | 1
| 1 | 2 | 362 | 358 | 348 | 353 | 357 | 390 | 375.9 | 355.8 | 353 | 356.78 | 385.6711 | 359.3 | 379 |  | 0
| 1 | 3 | 329 | 398 | 380 | 356 | 347 | 398.2341 | 364.1 | 350.1 | 359 | 345.53 | 381.8297 | 358.6 | 382 |  | 16
| 1 | 4 | 347 | 379 | 367 | 354 | 328 | 388.4719 | 364.9 | 354 | 356 | 346.49 | 385.1031 | 361.8 | 370 |  | 361.9022
| 1 | 5 | 332 | 361 | 373 | 347 | 334 | 385.53 | 374.4 | 356 | 350 | 345.59 | 385.275 | 354.6 | 373 |  | 106
| 1 | 6 | 319 | 378 | 349 | 363 | 336 | 389.1784 | 366.8 | 350.1 | 355 | 352.01 | 386.1991 | 362.2 | 373
| Mean |  | 339.6667 | 373.3333 | 364 | 355 | 339 | 390.7983 | 369.9167 | 352.7 | 354.1667 | 347.9283 | 384.7518 | 359.6333 | 373.8333
| Median |  | 339.5 | 372 | 367 | 355 | 335 | 389.5892 | 370.1 | 352.1 | 354 | 346.04 | 385.1891 | 360.3 | 373
| Std Dev. |  | 15.7438 | 14.8549 | 12.9306 | 5.2536 | 10.8812 | 4.4353 | 5.2297 | 2.8969 | 3.1885 | 5.5493 | 1.5477 | 2.8486 | 5.8452
| Rel.Std.Dev. |  | 4.635% | 3.979% | 3.552% | 1.48% | 3.21% | 1.135% | 1.414% | 0.8213% | 0.9003% | 1.595% | 0.4023% | 0.7921% | 1.564%
| PDM3 |  | -6.144% | 3.159% | 0.5797% | -1.907% | -6.328% | 7.984% | 2.215% | -2.543% | -2.137% | -3.861% | 6.314% | -0.6269% | 3.297%
| Z-Score (Absolute) |  | 1.15 | 0.79 | 0.25 | 0.27 | 1.19 | 1.8 | 0.59 | 0.4 | 0.31 | 0.67 | 1.45 | 0 | 0.82
| Analytical results for Sn in OREAS 462 (Certified Value 23.9 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 22.6 | 25.2 | 32.1 | 24.9 | 23.4 | 22.4183 | 24.3 | 22 | 19 | 24.8 | 24 |  |  |  | 1
| 1 | 2 | 22.4 | 25.1 | 28.1 | 24.6 | 25.1 | 24.0144 | 24.7 | 20 | 18.9 | 24.3 | 25 |  |  |  | 8
| 1 | 3 | 22.7 | 28 | 29.9 | 25.4 | 24 | 24.0934 | 24.3 | 23 | 19.2 | 24.9 | 25 |  |  |  | 16
| 1 | 4 | 21.7 | 25.6 | 29.5 | 24.9 | 23.6 | 23.0337 | 24 | 21 | 19.2 | 23.8 | 29 |  |  |  | 23.8764
| 1 | 5 | 21.6 | 25.8 | 29.4 | 24.7 | 23.4 | 21.5856 | 24.6 | 22 | 19 | 23.6 | 24 |  |  |  | 107
| 1 | 6 | 22.5 | 26 | 27.8 | 25.1 | 24.2 | 22.8413 | 24.5 | 22 | 18.6 | 24.7 | 26
| Mean |  | 22.25 | 25.95 | 29.4667 | 24.9333 | 23.95 | 22.9978 | 24.4 | 21.6667 | 18.9833 | 24.35 | 25.5
| Median |  | 22.45 | 25.7 | 29.45 | 24.9 | 23.8 | 22.9375 | 24.4 | 22 | 19 | 24.5 | 25
| Std Dev. |  | 0.47645 | 1.0616 | 1.5345 | 0.28752 | 0.65038 | 0.95811 | 0.25298 | 1.0328 | 0.22286 | 0.54681 | 1.8708
| Rel.Std.Dev. |  | 2.141% | 4.091% | 5.208% | 1.153% | 2.716% | 4.166% | 1.037% | 4.767% | 1.174% | 2.246% | 7.337%
| PDM3 |  | -6.812% | 8.685% | 23.41% | 4.427% | 0.3082% | -3.68% | 2.193% | -9.255% | -20.49% | 1.983% | 6.8%
| Z-Score (Absolute) |  | 1.19 | 0.67 | 2.9 | 0.33 | 0.23 | 0.77 | 0.03 | 1.52 | 3.04 | 0 | 0.25
| Analytical results for Sr in OREAS 462 (Certified Value 738 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 810 | 805 | 626 | 687 | 598 | 752.2599 | 755.9 | 795.1 | 747 | 762.7 | 762.1197 | 506 | 780 | 677 | 1
| 1 | 2 | 828 | 810 | 631 | 710 | 616 | 749.65 | 743 | 797.9 | 737 | 768.33 | 762.0783 | 500 | 770 | 713 | 10
| 1 | 3 | 782 | 824 | 634 | 710 | 618 | 744 | 734.7 | 798.2 | 749 | 771.18 | 762.3057 | 511 | 760 | 695 | 16
| 1 | 4 | 793 | 800 | 640 | 705 | 602 | 746.9309 | 746.7 | 806.9 | 748 | 803.24 | 762.3315 | 518 | 780 | 697 | 737.5233
| 1 | 5 | 806 | 797 | 637 | 687 | 684 | 752.7642 | 743 | 793.3 | 734 | 773.14 | 762.9623 | 512 | 770 | 713 | 108
| 1 | 6 | 775 | 807 | 643 | 700 | 615 | 744.5491 | 745.4 | 791.9 | 748 | 762.69 | 764.4186 | 509 | 775 | 704
| Mean |  | 799 | 807.1667 | 635.1667 | 699.8333 | 622.1667 | 748.359 | 744.7833 | 797.2167 | 743.8333 | 773.5467 | 762.7027 | 509.3333 | 772.5 | 699.8333
| Median |  | 799.5 | 806 | 635.5 | 702.5 | 615.5 | 748.2904 | 744.2 | 796.5 | 747.5 | 769.755 | 762.3186 | 510 | 772.5 | 700.5
| Std Dev. |  | 19.5551 | 9.4956 | 6.1779 | 10.6097 | 31.3714 | 3.7907 | 6.8654 | 5.3533 | 6.5549 | 15.167 | 0.89853 | 6.0553 | 7.5829 | 13.5413
| Rel.Std.Dev. |  | 2.447% | 1.176% | 0.9726% | 1.516% | 5.042% | 0.5065% | 0.9218% | 0.6715% | 0.8812% | 1.961% | 0.1178% | 1.189% | 0.9816% | 1.935%
| PDM3 |  | 8.336% | 9.443% | -13.88% | -5.11% | -15.64% | 1.469% | 0.9844% | 8.094% | 0.8556% | 4.884% | 3.414% | -30.94% | 4.742% | -5.11%
| Z-Score (Absolute) |  | 0.76 | 0.87 | 1.61 | 0.67 | 1.97 | 0.03 | 0.03 | 0.73 | 0.04 | 0.3 | 0.23 | 3.42 | 0.37 | 0.67
| Analytical results for Ta in OREAS 462 (Certified Value 21.8 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 24.3 | 24.5 | 23.2 | 22.8 | 20.3426 | 22 | 13.7 | 11.6 | 18 | 20.6 |  |  |  |  | 1
| 1 | 2 | 24.5 | 25.2 | 22.8 | 24.4 | 19.5239 | 21.8 | 11.9 | 11.8 | 15.21 | 21.2 |  |  |  |  | 11
| 1 | 3 | 23.8 | 25.7 | 23.5 | 23.1 | 19.2104 | 21.8 | 12.9 | 11.2 | 16.97 | 20.9 |  |  |  |  | 16
| 1 | 4 | 23.7 | 25.4 | 23.1 | 22.6 | 20.0208 | 21.1 | 11.8 | 11.3 | 16.78 | 21.3 |  |  |  |  | 21.8115
| 1 | 5 | 23.7 | 25.7 | 22.6 | 22.5 | 20.1645 | 21.2 | 11.9 | 10.9 | 16.87 | 20.5 |  |  |  |  | 109
| 1 | 6 | 23.2 | 24.6 | 23.4 | 23.2 | 19.6495 | 21.3 | 12.5 | 11.6 | 16.88 | 21.8
| Mean |  | 23.8667 | 25.1833 | 23.1 | 23.1 | 19.8186 | 21.5333 | 12.45 | 11.4 | 16.785 | 21.05
| Median |  | 23.75 | 25.3 | 23.15 | 22.95 | 19.8352 | 21.55 | 12.2 | 11.45 | 16.875 | 21.05
| Std Dev. |  | 0.46762 | 0.52694 | 0.34641 | 0.69282 | 0.42903 | 0.37771 | 0.74766 | 0.32863 | 0.89525 | 0.48477
| Rel.Std.Dev. |  | 1.959% | 2.092% | 1.5% | 2.999% | 2.165% | 1.754% | 6.005% | 2.883% | 5.334% | 2.303%
| PDM3 |  | 9.422% | 15.46% | 5.907% | 5.907% | -9.137% | -1.275% | -42.92% | -47.73% | -23.05% | -3.491%
| Z-Score (Absolute) |  | 0.79 | 1.2 | 0.56 | 0.48 | 0.45 | 0.07 | 2.72 | 3.04 | 1.29 | 0.07
| Analytical results for Tb in OREAS 462 (Certified Value 14.7 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 15.35 | 14.5 | 15.9 | 13.85 | 15.45 | 15.093 | 14.7 | 14.4 | 17.6 | 14.58 | 21.1375 | 14.25 | 12.5 |  | 1
| 1 | 2 | 15.5 | 14.45 | 15.65 | 13.9 | 16.65 | 15.0849 | 14.5 | 14.59 | 17.6 | 15.17 | 21.9763 | 14.25 | 12.8 |  | 1
| 1 | 3 | 13.85 | 15.8 | 16.2 | 13.75 | 15.5 | 14.9999 | 14.5 | 14.32 | 17.4 | 14.85 | 22.2116 | 14.2 | 13.1 |  | 16
| 1 | 4 | 15.35 | 14.4 | 14.95 | 13.8 | 15.1 | 14.6001 | 14.7 | 14.9 | 17.4 | 14.63 | 21.9933 | 14.45 | 12.9 |  | 14.7025
| 1 | 5 | 14.4 | 13.85 | 15.6 | 13.2 | 15.55 | 14.74 | 14.6 | 14.46 | 17.4 | 14.67 | 22.2473 | 14 | 12.8 |  | 110
| 1 | 6 | 13.7 | 15.1 | 14.7 | 13.8 | 15.4 | 14.9192 | 14.5 | 14.58 | 17.6 | 14.66 | 20.2038 | 14.3 | 13
| Mean |  | 14.6917 | 14.6833 | 15.5 | 13.7167 | 15.6083 | 14.9062 | 14.5833 | 14.5417 | 17.5 | 14.76 | 21.6283 | 14.2417 | 12.85
| Median |  | 14.875 | 14.475 | 15.625 | 13.8 | 15.475 | 14.9595 | 14.55 | 14.52 | 17.5 | 14.665 | 21.9848 | 14.25 | 12.85
| Std Dev. |  | 0.81204 | 0.67577 | 0.57009 | 0.2582 | 0.53424 | 0.19856 | 0.098319 | 0.204 | 0.10954 | 0.22073 | 0.80595 | 0.14634 | 0.20736
| Rel.Std.Dev. |  | 5.527% | 4.602% | 3.678% | 1.882% | 3.423% | 1.332% | 0.6742% | 1.403% | 0.626% | 1.495% | 3.726% | 1.028% | 1.614%
| PDM3 |  | -0.07335% | -0.13% | 5.425% | -6.705% | 6.161% | 1.386% | -0.8102% | -1.094% | 19.03% | 0.3914% | 47.11% | -3.134% | -12.6%
| Z-Score (Absolute) |  | 0 | 0.01 | 1.21 | 1.46 | 1.06 | 0.32 | 0.16 | 0.22 | 4.21 | 0.1 | 10.82 | 0.67 | 2.76
| Analytical results for Te in OREAS 462 (Certified Value 0.33 ppm) | [15 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.36 | 0.37 | 0.52 | 0.31 | 0.3 | 0.22279 | < 1 | 0.6 | 0.47 | 0.3 | < 0.2 |  |  |  | 1
| 1 | 2 | 0.3 | 0.39 | 0.47 | 0.4 | 0.29 | 0.23911 | < 1 | < 0.5 | 0.4 | 0.3 | 0.2 |  |  |  | 16
| 1 | 3 | 0.31 | 0.42 | 0.53 | 0.39 | 0.31 | 0.20702 | < 1 | 0.5 | 0.41 | 0.3 | 0.2 |  |  |  | 16
| 1 | 4 | 0.33 | 0.38 | 0.53 | 0.33 | 0.3 | 0.27671 | < 1 | < 0.5 | 0.41 | 0.3 | 0.2 |  |  |  | 0.33312
| 1 | 5 | 0.32 | 0.4 | 0.55 | 0.31 | 0.3 | 0.24 | < 1 | < 0.5 | 0.44 | 0.3 | 0.2 |  |  |  | 111
| 1 | 6 | 0.32 | 0.38 | 0.47 | 0.35 | 0.32 | 0.26535 | < 1 | < 0.5 | 0.42 | 0.3 | 0.2
| Mean |  | 0.32333 | 0.39 | 0.51167 | 0.34833 | 0.30333 | 0.24183 |  | 0.55 | 0.425 | 0.3 | 0.2
| Median |  | 0.32 | 0.385 | 0.525 | 0.34 | 0.3 | 0.23956 |  | 0.55 | 0.415 | 0.3 | 0.2
| Std Dev. |  | 0.020656 | 0.017889 | 0.033714 | 0.0392 | 0.010328 | 0.025905 |  | 0.070711 | 0.025884 | 0 | 0
| Rel.Std.Dev. |  | 6.388% | 4.587% | 6.589% | 11.25% | 3.405% | 10.71% |  | 12.86% | 6.09% | 0% | 0%
| PDM3 |  | -2.938% | 17.08% | 53.6% | 4.567% | -8.941% | -27.4% |  | 65.11% | 27.58% | -9.942% | -39.96%
| Z-Score (Absolute) |  | 0.12 | 0.51 | 1.65 | 0.12 | 0.31 | 0.88 | NA | 2.02 | 0.84 | 0.34 | 1.28
| Analytical results for Th in OREAS 462 (Certified Value 246 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 284 | 288 | 275 | 241 | 248 | 254.9739 | 248.1 | 250 | 235.21 | 265.5161 | 239 |  |  |  | 1
| 1 | 2 | 285 | 278 | 263 | 242 | 264 | 252.1338 | 244.8 | 245 | 243.47 | 266.0494 | 246 |  |  |  | 2
| 1 | 3 | 225 | 296 | 278 | 244 | 252 | 252.7165 | 248.9 | 250 | 239.45 | 264.7745 | 253 |  |  |  | 16
| 1 | 4 | 260 | 285 | 221 | 243 | 247 | 241.2511 | 248.2 | 254 | 230.54 | 265.4295 | 243 |  |  |  | 245.568
| 1 | 5 | 244 | 287 | 276 | 234 | 257 | 237.2 | 250.4 | 249 | 237.73 | 266.8418 | 247 |  |  |  | 112
| 1 | 6 | 224 | 284 | 269 | 246 | 250 | 244.5063 | 250 | 252 | 242.98 | 265.8121 | 250
| Mean |  | 253.6667 | 286.3333 | 263.6667 | 241.6667 | 253 | 247.1303 | 248.4 | 250 | 238.23 | 265.7372 | 246.3333
| Median |  | 252 | 286 | 272 | 242.5 | 251 | 248.3201 | 248.55 | 250 | 238.59 | 265.6641 | 246.5
| Std Dev. |  | 27.3398 | 5.8878 | 21.6117 | 4.1312 | 6.4498 | 7.1808 | 1.995 | 3.0332 | 4.8993 | 0.6919 | 4.9666
| Rel.Std.Dev. |  | 10.78% | 2.056% | 8.197% | 1.709% | 2.549% | 2.906% | 0.8031% | 1.213% | 2.057% | 0.2604% | 2.016%
| PDM3 |  | 3.298% | 16.6% | 7.37% | -1.589% | 3.026% | 0.6362% | 1.153% | 1.805% | -2.988% | 8.213% | 0.3117%
| Z-Score (Absolute) |  | 1.32 | 4.92 | 3.09 | 0.67 | 0.6 | 0.16 | 0 | 0.21 | 1.32 | 2.25 | 0.27
| Analytical results for Ti in OREAS 462 (Certified Value 0.986 wt.%) | [15 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  | 3
3
| 1 | 1 | 1.15 | 1.2 | 0.995 | 0.883 | 0.882 | 0.98067 | 0.94 | 0.7302 | 0.862 | 1.2365 | 0.68 | 1.34 |  |  | 1
| 1 | 2 | 1.14 | 1.145 | 1.005 | 0.941 | 0.884 | 0.98384 | 0.94 | 0.6595 | 0.881 | 0.9452 | 0.64 | 1.36 |  |  | 3
| 1 | 3 | 1.09 | 1.22 | 0.997 | 0.976 | 0.872 | 0.9388 | 0.91 | 0.6702 | 0.856 | 0.9618 | 0.66 | 1.32 |  |  | 16
| 1 | 4 | 1.095 | 1.19 | 1.02 | 0.952 | 0.907 | 0.99168 | 0.89 | 0.6619 | 0.864 | 0.9535 | 0.65 | 1.38 |  |  | 0.98596
| 1 | 5 | 1.11 | 1.165 | 0.995 | 0.918 | 0.956 | 0.96969 | 0.91 | 0.7028 | 0.857 | 0.9549 | 0.7 | 1.35 |  |  | 113
| 1 | 6 | 1.105 | 1.15 | 0.983 | 0.897 | 0.942 | 0.93244 | 0.89 | 0.718 | 0.888 | 1.1778 | 0.67 | 1.43
| Mean |  | 1.115 | 1.1783 | 0.99917 | 0.92783 | 0.90717 | 0.96619 | 0.91333 | 0.69043 | 0.868 | 1.0383 | 0.66667 | 1.3633
| Median |  | 1.1075 | 1.1775 | 0.996 | 0.9295 | 0.8955 | 0.97518 | 0.91 | 0.6865 | 0.863 | 0.95835 | 0.665 | 1.355
| Std Dev. |  | 0.024495 | 0.029777 | 0.0124 | 0.035028 | 0.034655 | 0.02479 | 0.022509 | 0.030577 | 0.013312 | 0.13222 | 0.021602 | 0.038297
| Rel.Std.Dev. |  | 2.197% | 2.527% | 1.241% | 3.775% | 3.82% | 2.566% | 2.465% | 4.429% | 1.534% | 12.73% | 3.24% | 2.809%
| PDM3 |  | 13.09% | 19.51% | 1.339% | -5.896% | -7.992% | -2.006% | -7.366% | -29.97% | -11.96% | 5.307% | -32.38% | 38.27%
| Z-Score (Absolute) |  | 1.73 | 2.38 | 0.54 | 0.2 | 0.41 | 0.2 | 0.35 | 2.64 | 0.81 | 0.53 | 2.88 | 4.28
| Analytical results for Tl in OREAS 462 (Certified Value 0.083 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.08 | 0.17 | 0.09 | 0.09 | 0.08 | 0.070339 | 0.07 | < 0.2 | 0.08 | 0.07 | 0.060455 | 0.1 | < 0.1 |  | 1
| 1 | 2 | 0.09 | 0.14 | 0.09 | 0.09 | 0.09 | 0.076687 | 0.06 | < 0.2 | 0.08 | 0.07 | 0.080606 | 0.1 | < 0.1 |  | 17
| 1 | 3 | 0.09 | 0.15 | 0.09 | 0.08 | 0.09 | 0.076876 | 0.08 | < 0.2 | 0.08 | 0.07 | 0.092276 | 0.1 | < 0.1 |  | 16
| 1 | 4 | 0.08 | 0.12 | 0.08 | 0.07 | 0.09 | 0.087556 | 0.09 | < 0.2 | 0.08 | 0.07 | 0.12091 | 0.1 | < 0.1 |  | 0.082847
| 1 | 5 | 0.08 | 0.1 | 0.09 | 0.08 | 0.09 | 0.084067 | 0.07 | < 0.2 | 0.08 | 0.06 | 0.15379 | 0.1 | < 0.1 |  | 114
| 1 | 6 | 0.08 | 0.12 | 0.09 | 0.09 | 0.08 | 0.088201 | 0.08 | < 0.2 | 0.08 | 0.07 | 0.16334 | 0.1 | < 0.1
| Mean |  | 0.083333 | 0.13333 | 0.088333 | 0.083333 | 0.086667 | 0.080621 | 0.075 |  | 0.08 | 0.068333 | 0.1119 | 0.1
| Median |  | 0.08 | 0.13 | 0.09 | 0.085 | 0.09 | 0.080471 | 0.075 |  | 0.08 | 0.07 | 0.10659 | 0.1
| Std Dev. |  | 0.005164 | 0.025033 | 0.0040825 | 0.008165 | 0.005164 | 0.0071083 | 0.010488 |  | 0 | 0.0040825 | 0.041213 | 1.5202e-17
| Rel.Std.Dev. |  | 6.197% | 18.77% | 4.622% | 9.798% | 5.958% | 8.817% | 13.98% |  | 0% | 5.974% | 36.83% | 1.52e-14%
| PDM3 |  | 0.5873% | 60.94% | 6.623% | 0.5873% | 4.611% | -2.687% | -9.471% |  | -3.436% | -17.52% | 35.06% | 20.7%
| Z-Score (Absolute) |  | 0 | 6.74 | 0.67 | 0 | 0.45 | 0.37 | 1.12 | NA | 0.45 | 2.02 | 3.85 | 2.25 | NA
| Analytical results for Tm in OREAS 462 (Certified Value 1.01 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 1.17 | 1.06 | 1.07 | 0.98 | 1.01 | 0.86151 | 0.97329 | 1 | 0.95 | 0.87 | 1.1972 | 1 | 0.88 |  | 1
| 1 | 2 | 1.17 | 1.05 | 1.05 | 0.98 | 1.1 | 0.993 | 0.9261 | 1 | 0.94 | 0.86 | 1.1919 | 1 | 0.94 |  | 4
| 1 | 3 | 1.07 | 1.17 | 1.15 | 0.99 | 1.06 | 0.99403 | 0.96935 | 1.1 | 0.98 | 0.85 | 1.2283 | 1 | 0.96 |  | 16
| 1 | 4 | 1.11 | 1.04 | 1.1 | 0.98 | 1.01 | 0.87853 | 0.98107 | 1 | 0.95 | 0.92 | 1.164 | 0.95 | 0.94 |  | 1.011
| 1 | 5 | 1.09 | 1.04 | 1.09 | 0.95 | 1.02 | 0.88455 | 0.9855 | 1.1 | 0.95 | 0.84 | 1.1914 | 0.95 | 0.94 |  | 115
| 1 | 6 | 1.06 | 1.1 | 1.05 | 0.98 | 1.03 | 0.927 | 1.0015 | 1.1 | 0.95 | 0.89 | 1.1245 | 1 | 0.96
| Mean |  | 1.1117 | 1.0767 | 1.085 | 0.97667 | 1.0383 | 0.9231 | 0.9728 | 1.05 | 0.95333 | 0.87167 | 1.1829 | 0.98333 | 0.93667
| Median |  | 1.1 | 1.055 | 1.08 | 0.98 | 1.025 | 0.90578 | 0.97718 | 1.05 | 0.95 | 0.865 | 1.1917 | 1 | 0.94
| Std Dev. |  | 0.048339 | 0.050859 | 0.037815 | 0.013663 | 0.035449 | 0.058648 | 0.025476 | 0.054772 | 0.013663 | 0.029269 | 0.035159 | 0.02582 | 0.029439
| Rel.Std.Dev. |  | 4.348% | 4.724% | 3.485% | 1.399% | 3.414% | 6.353% | 2.619% | 5.216% | 1.433% | 3.358% | 2.972% | 2.626% | 3.143%
| PDM3 |  | 9.953% | 6.491% | 7.316% | -3.399% | 2.7% | -8.697% | -3.782% | 3.854% | -5.707% | -13.78% | 17% | -2.74% | -7.356%
| Z-Score (Absolute) |  | 1.44 | 0.84 | 1.14 | 0.07 | 0.62 | 0.67 | 0.12 | 0.75 | 0.34 | 1.25 | 2.23 | 0 | 0.52
| Analytical results for U in OREAS 462 (Certified Value 5.79 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 5.9 | 5.7 | 6.4 | 5.8 | 6 | 5.6058 | 6.05 | 5.7 | 6.01 | 5.45 | 6.7447 | 5.7 | 5 |  | 1
| 1 | 2 | 5.9 | 5.9 | 6 | 5.8 | 6.4 | 5.4412 | 5.95 | 5.5 | 6.03 | 5.56 | 6.858 | 5.7 | 5.2 |  | 5
| 1 | 3 | 5.8 | 6.4 | 6.4 | 5.8 | 6.1 | 5.4339 | 5.79 | 5.6 | 5.83 | 6.9 | 6.7624 | 5.5 | 5.1 |  | 16
| 1 | 4 | 5.8 | 5.8 | 4.7 | 5.8 | 5.9 | 5.6513 | 5.86 | 5.5 | 5.98 | 8.29 | 6.6925 | 5.7 | 5 |  | 5.7922
| 1 | 5 | 5.7 | 5.8 | 6.6 | 5.6 | 6 | 5.4565 | 5.95 | 5.8 | 5.81 | 6.87 | 7.0926 | 5.6 | 5.1 |  | 116
| 1 | 6 | 5.6 | 6.1 | 6.2 | 5.9 | 6 | 5.5642 | 5.83 | 5.7 | 5.87 | 5.51 | 6.3265 | 5.6 | 5.1
| Mean |  | 5.7833 | 5.95 | 6.05 | 5.7833 | 6.0667 | 5.5255 | 5.905 | 5.6333 | 5.9217 | 6.43 | 6.7461 | 5.6333 | 5.0833
| Median |  | 5.8 | 5.85 | 6.3 | 5.8 | 6 | 5.5104 | 5.905 | 5.65 | 5.925 | 6.215 | 6.7536 | 5.65 | 5.1
| Std Dev. |  | 0.1169 | 0.25884 | 0.6921 | 0.098319 | 0.17512 | 0.093858 | 0.095864 | 0.12111 | 0.096419 | 1.1347 | 0.24971 | 0.08165 | 0.075277
| Rel.Std.Dev. |  | 2.021% | 4.35% | 11.44% | 1.7% | 2.887% | 1.699% | 1.623% | 2.15% | 1.628% | 17.65% | 3.702% | 1.449% | 1.481%
| PDM3 |  | -0.1537% | 2.724% | 4.45% | -0.1537% | 4.738% | -4.606% | 1.947% | -2.743% | 2.235% | 11.01% | 16.47% | -2.743% | -12.24%
| Z-Score (Absolute) |  | 0.26 | 0 | 1.57 | 0.26 | 0.48 | 1.14 | 0.15 | 0.77 | 0.21 | 0.67 | 3.02 | 0.77 | 2.65
| Analytical results for V in OREAS 462 (Certified Value 353 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 329 | 371 | 270 | 286 | 265 | 350.0606 | 360.4 | 375 | 339 | 390 | 357.63 | 162 | 370 |  | 1
| 1 | 2 | 327 | 371 | 271 | 297 | 268 | 358.6653 | 356.1 | 367 | 341 | 406 | 357.297 | 161 | 365 |  | 6
| 1 | 3 | 322 | 379 | 271 | 302 | 272 | 340.5233 | 360.1 | 366 | 343 | 388 | 356.812 | 162 | 360 |  | 16
| 1 | 4 | 321 | 365 | 273 | 301 | 270 | 354.4234 | 360 | 359 | 355 | 395 | 357.541 | 160 | 360 |  | 352.6418
| 1 | 5 | 325 | 364 | 269 | 293 | 296 | 335.3191 | 361.4 | 364 | 349 | 392 | 358.174 | 162 | 370 |  | 117
| 1 | 6 | 320 | 368 | 273 | 295 | 278 | 357.8372 | 361.4 | 362 | 351 | 383 | 358.823 | 157 | 370
| Mean |  | 324 | 369.6667 | 271.1667 | 295.6667 | 274.8333 | 349.4715 | 359.9 | 365.5 | 346.3333 | 392.3333 | 357.7128 | 160.6667 | 365.8333
| Median |  | 323.5 | 369.5 | 271 | 296 | 271 | 352.242 | 360.25 | 365 | 346 | 391 | 357.5855 | 161.5 | 367.5
| Std Dev. |  | 3.5777 | 5.4283 | 1.6021 | 5.8538 | 11.2502 | 9.589 | 1.9616 | 5.4681 | 6.2823 | 7.8145 | 0.70227 | 1.9664 | 4.916
| Rel.Std.Dev. |  | 1.104% | 1.468% | 0.5908% | 1.98% | 4.093% | 2.744% | 0.545% | 1.496% | 1.814% | 1.992% | 0.1963% | 1.224% | 1.344%
| PDM3 |  | -8.122% | 4.828% | -23.1% | -16.16% | -22.06% | -0.899% | 2.058% | 3.646% | -1.789% | 11.26% | 1.438% | -54.44% | 3.741%
| Z-Score (Absolute) |  | 0.85 | 0.67 | 2.61 | 1.8 | 2.63 | 0 | 0.35 | 0.54 | 0.1 | 1.43 | 0.28 | 6.3 | 0.55
| Analytical results for W in OREAS 462 (Certified Value 2.31 ppm) | [15 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.7 | 2.6 | 3.4 | 2.7 | 2.7 | 2.1439 | 2 | 2 | 2.2 | 2.2 | 2 |  |  |  | 1
| 1 | 2 | 2.8 | 2.6 | 3.2 | 2.7 | 2.9 | 2.1152 | 2 | 2 | 2.2 | 1.7 | 2.5 |  |  |  | 7
| 1 | 3 | 2.7 | 2.9 | 3.4 | 2.8 | 2.7 | 2.4708 | 2 | 2 | 2 | 1.8 | 2 |  |  |  | 16
| 1 | 4 | 2.7 | 2.6 | 3.3 | 2.7 | 2.7 | 2.4337 | 2 | 2 | 2 | 1.8 | 2 |  |  |  | 2.3069
| 1 | 5 | 2.6 | 2.6 | 3.3 | 2.6 | 2.6 | 2.1658 | 2 | 2 | 2.4 | 1.9 | 2 |  |  |  | 118
| 1 | 6 | 2.6 | 2.6 | 3.2 | 2.7 | 2.7 | 2.2821 | 2 | 2 | 2.1 | 1.9 | 2.5
| Mean |  | 2.6833 | 2.65 | 3.3 | 2.7 | 2.7167 | 2.2686 | 2 | 2 | 2.15 | 1.8833 | 2.1667
| Median |  | 2.7 | 2.6 | 3.3 | 2.7 | 2.7 | 2.2239 | 2 | 2 | 2.15 | 1.85 | 2
| Std Dev. |  | 0.075277 | 0.12247 | 0.089443 | 0.063246 | 0.098319 | 0.1536 | 0 | 0 | 0.15166 | 0.17224 | 0.2582
| Rel.Std.Dev. |  | 2.805% | 4.622% | 2.71% | 2.342% | 3.619% | 6.771% | 0% | 0% | 7.054% | 9.145% | 11.92%
| PDM3 |  | 16.32% | 14.87% | 43.05% | 17.04% | 17.76% | -1.659% | -13.3% | -13.3% | -6.8% | -18.36% | -6.077%
| Z-Score (Absolute) |  | 0.84 | 0.67 | 2.1 | 0.88 | 0.84 | 0 | 0.55 | 0.55 | 0.24 | 0.91 | 0.21
| Analytical results for Y in OREAS 462 (Certified Value 125 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 130.5 | 115.5 | 138.5 | 66.7 | 126 | 124.7308 | 130.4 | 126.8 | 130 | 128.83 | 152.3358 | 117.7 | 121 |  | 1
| 1 | 2 | 133 | 119 | 135 | 65.9 | 136.5 | 129.0619 | 128.7 | 126.1 | 127 | 126.96 | 155.1019 | 119.1 | 128 |  | 8
| 1 | 3 | 117.5 | 124.5 | 141.5 | 66 | 132.5 | 127.0213 | 127.7 | 126.7 | 130 | 127.95 | 155.0102 | 119.3 | 124 |  | 16
| 1 | 4 | 121 | 119 | 140.5 | 65.6 | 128.5 | 123.3163 | 126.2 | 127.1 | 131 | 124.94 | 152.4945 | 119.5 | 129 |  | 125.2743
| 1 | 5 | 120 | 119 | 146 | 64.4 | 130.5 | 120.9701 | 127.3 | 125.9 | 127 | 126.69 | 156.0121 | 118.4 | 129 |  | 119
| 1 | 6 | 117.5 | 118 | 133 | 67.5 | 129.5 | 125.2079 | 127 | 126.2 | 130 | 130.08 | 153.2525 | 118 | 125
| Mean |  | 123.25 | 119.1667 | 139.0833 | 66.0167 | 130.5833 | 125.0514 | 127.8833 | 126.4667 | 129.1667 | 127.575 | 154.0345 | 118.6667 | 126
| Median |  | 120.5 | 119 | 139.5 | 65.95 | 130 | 124.9694 | 127.5 | 126.45 | 130 | 127.455 | 154.1314 | 118.75 | 126.5
| Std Dev. |  | 6.7731 | 2.9439 | 4.6842 | 1.0458 | 3.6113 | 2.8199 | 1.4825 | 0.46762 | 1.7224 | 1.7942 | 1.5408 | 0.73937 | 3.2249
| Rel.Std.Dev. |  | 5.495% | 2.47% | 3.368% | 1.584% | 2.766% | 2.255% | 1.159% | 0.3698% | 1.333% | 1.406% | 1% | 0.6231% | 2.559%
| PDM3 |  | -1.616% | -4.875% | 11.02% | -47.3% | 4.238% | -0.1779% | 2.083% | 0.9518% | 3.107% | 1.837% | 22.96% | -5.275% | 0.5793%
| Z-Score (Absolute) |  | 0.67 | 1.75 | 2.64 | 12.67 | 0.86 | 0.3 | 0.3 | 0 | 0.57 | 0.23 | 5.78 | 1.64 | 0.1
| Analytical results for Yb in OREAS 462 (Certified Value 4.36 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.69 | 4.26 | 4.38 | 4.1 | 4.22 | 4.0744 | 4.2591 | 4.6 | 5.52 | 3.88 | 5.3172 | 4.5 | 4.35 |  | 1
| 1 | 2 | 4.66 | 4.38 | 4.23 | 4.22 | 4.53 | 3.9469 | 4.1193 | 4.2 | 5.43 | 3.72 | 5.3544 | 4.55 | 4.6 |  | 9
| 1 | 3 | 4.29 | 4.7 | 4.58 | 4.1 | 4.33 | 3.8374 | 4.0998 | 4.3 | 5.4 | 3.7 | 5.2194 | 4.4 | 4.85 |  | 16
| 1 | 4 | 4.4 | 4.21 | 4.47 | 4.07 | 4.15 | 4.0994 | 3.9212 | 4.2 | 5.41 | 3.77 | 5.1273 | 4.5 | 4.9 |  | 4.3644
| 1 | 5 | 4.38 | 4.27 | 4.52 | 4.03 | 4.26 | 4.18 | 4.1822 | 4.5 | 5.4 | 3.96 | 5.5412 | 4.55 | 4.55 |  | 120
| 1 | 6 | 4.33 | 4.42 | 4.61 | 4.14 | 4.3 | 3.908 | 3.9279 | 4.2 | 5.88 | 4.19 | 4.9215 | 4.55 | 4.8
| Mean |  | 4.4583 | 4.3733 | 4.465 | 4.11 | 4.2983 | 4.0077 | 4.0849 | 4.3333 | 5.5067 | 3.87 | 5.2468 | 4.5083 | 4.675
| Median |  | 4.39 | 4.325 | 4.495 | 4.1 | 4.28 | 4.0107 | 4.1096 | 4.25 | 5.42 | 3.825 | 5.2683 | 4.525 | 4.7
| Std Dev. |  | 0.17244 | 0.1784 | 0.14124 | 0.065115 | 0.12983 | 0.13053 | 0.13614 | 0.17512 | 0.18843 | 0.18547 | 0.21167 | 0.058452 | 0.21154
| Rel.Std.Dev. |  | 3.868% | 4.079% | 3.163% | 1.584% | 3.021% | 3.257% | 3.333% | 4.041% | 3.422% | 4.793% | 4.034% | 1.297% | 4.525%
| PDM3 |  | 2.153% | 0.2057% | 2.306% | -5.828% | -1.513% | -8.172% | -6.402% | -0.7108% | 26.17% | -11.33% | 20.22% | 3.299% | 7.118%
| Z-Score (Absolute) |  | 0.22 | 0 | 0.23 | 0.67 | 0.19 | 0.94 | 0.74 | 0.1 | 2.71 | 1.29 | 2.09 | 0.35 | 0.77
| Analytical results for Zn in OREAS 462 (Certified Value 244 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 240 | 259 | 193 | 191 | 186 | 253.6595 | 255.0775 | 287 | 262 | 275 | 318.851 | 158 | 250 | 244 | 1
| 1 | 2 | 240 | 255 | 193 | 201 | 187 | 253.4656 | 249.9438 | 258 | 266 | 275 | 319.047 | 159 | 250 | 246 | 0
| 1 | 3 | 233 | 261 | 192 | 204 | 187 | 251.2351 | 248.6723 | 269 | 263 | 262 | 319.312 | 163 | 250 | 252 | 16
| 1 | 4 | 236 | 254 | 194 | 201 | 189 | 257.162 | 250.7766 | 259 | 265 | 271 | 318.015 | 162 | 250 | 246 | 243.663
| 1 | 5 | 237 | 249 | 192 | 196 | 202 | 257.1701 | 250.3992 | 266 | 259 | 266 | 315.008 | 160 | 250 | 262 | 121
| 1 | 6 | 235 | 251 | 193 | 197 | 194 | 255.379 | 248.0182 | 259 | 262 | 264 | 318.333 | 154 | 255 | 250
| Mean |  | 236.8333 | 254.8333 | 192.8333 | 198.3333 | 190.8333 | 254.6786 | 250.4813 | 266.3333 | 262.8333 | 268.8333 | 318.0943 | 159.3333 | 250.8333 | 250
| Median |  | 236.5 | 254.5 | 193 | 199 | 188 | 254.5193 | 250.1715 | 262.5 | 262.5 | 268.5 | 318.592 | 159.5 | 250 | 248
| Std Dev. |  | 2.7869 | 4.5789 | 0.75277 | 4.6332 | 6.1779 | 2.3338 | 2.4838 | 11.0574 | 2.4833 | 5.6362 | 1.5844 | 3.2042 | 2.0412 | 6.5727
| Rel.Std.Dev. |  | 1.177% | 1.797% | 0.3904% | 2.336% | 3.237% | 0.9164% | 0.9916% | 4.152% | 0.9448% | 2.097% | 0.4981% | 2.011% | 0.8138% | 2.629%
| PDM3 |  | -2.803% | 4.584% | -20.86% | -18.6% | -21.68% | 4.521% | 2.798% | 9.304% | 7.868% | 10.33% | 30.55% | -34.61% | 2.943% | 2.601%
| Z-Score (Absolute) |  | 0.72 | 0.22 | 3 | 2.71 | 3.22 | 0.21 | 0.01 | 0.6 | 0.63 | 0.94 | 3.5 | 4.74 | 0.01 | 0.16
| Analytical results for Zr in OREAS 462 (Certified Value 270 ppm) | [15 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 293 | 279 | 305 | 290 | 252 | 248.8969 | 221 | 338 | 238.7 | 124 | 260 | 281 |  |  | 1
| 1 | 2 | 290 | 279 | 299 | 287 | 262 | 253.3534 | 213 | 326 | 214.9 | 129 | 260 | 278 |  |  | 10
| 1 | 3 | 278 | 294 | 311 | 291 | 256 | 250.5112 | 218 | 318 | 226.2 | 128 | 255 | 271 |  |  | 16
| 1 | 4 | 279 | 289 | 314 | 288 | 255 | 244.9149 | 217 | 318 | 227 | 117 | 260 | 294 |  |  | 270.2672
| 1 | 5 | 285 | 278 | 321 | 280 | 269 | 250.6507 | 221 | 304 | 232.2 | 123 | 255 | 278 |  |  | 122
| 1 | 6 | 284 | 289 | 295 | 287 | 262 | 242.6079 | 220 | 308 | 232.3 | 117 | 245 | 279
| Mean |  | 284.8333 | 284.6667 | 307.5 | 287.1667 | 259.3333 | 248.4892 | 218.3333 | 318.6667 | 228.55 | 123 | 255.8333 | 280.1667
| Median |  | 284.5 | 284 | 308 | 287.5 | 259 | 249.7041 | 219 | 318 | 229.6 | 123.5 | 257.5 | 278.5
| Std Dev. |  | 5.9133 | 6.8313 | 9.7108 | 3.8687 | 6.186 | 3.9988 | 3.0768 | 12.3072 | 8.058 | 5.1769 | 5.8452 | 7.5741
| Rel.Std.Dev. |  | 2.076% | 2.4% | 3.158% | 1.347% | 2.385% | 1.609% | 1.409% | 3.862% | 3.526% | 4.209% | 2.285% | 2.703%
| PDM3 |  | 5.39% | 5.328% | 13.78% | 6.253% | -4.046% | -8.058% | -19.22% | 17.91% | -15.44% | -54.49% | -5.341% | 3.663%
| Z-Score (Absolute) |  | 0.57 | 0.57 | 1.36 | 0.66 | 0.31 | 0.69 | 1.74 | 1.75 | 1.39 | 5.07 | 0.36 | 0.31
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.