## Abbreviations
Table 1. Abbreviations used for OREAS 464
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 464
Laboratory Name and Location
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
Nagrom, Perth, WA, Australia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
SGS Vostok Limited, Chita, Russian Federation
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 464
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate Fusion XRF
Ce, Cerium (wt.%) | 1.5426 | 0.022149 | 1.5207 | 1.5646 | 1.5241 | 1.5612
Fe, Iron (wt.%) | 38.0367 | 0.72116 | 37.063 | 39.0103 | 37.7265 | 38.3468
La, Lanthanum (wt.%) | 1.1991 | 0.01683 | 1.1848 | 1.2135 | 1.1807 | 1.2176
Nd, Neodymium (wt.%) | 0.96367 | 0.030982 | 0.93013 | 0.99722 | 0.94463 | 0.98272
Pr, Praseodymium (ppm) | 2540.6949 | 81.564 | 2472.4749 | 2608.9149 | 2406.4373 | 2674.9525
Sm, Samarium (ppm) | 1409.7243 | 47.6299 | 1365.95 | 1453.4986 | IND | IND
Th, Thorium (ppm) | 506.0026 | 74.2303 | 432.0105 | 579.9948 | IND | IND
Y, Yttrium (ppm) | 442.4618 | 54.1764 | 394.9281 | 489.9955 | IND | IND
Thermogravimetry
LOI, Loss On Ignition @1000°C (wt.%) | 0.6969 | 0.0711 | 0.60808 | 0.78572 | 0.65228 | 0.74152
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 4.6305 | 0.15029 | 4.5179 | 4.743 | 4.5337 | 4.7272
Ba, Barium (ppm) | 1656.7244 | 69.2874 | 1614.0407 | 1699.4081 | 1612.9155 | 1700.5332
Be, Beryllium (ppm) | 15.6289 | 1.5763 | 14.1131 | 17.1446 | IND | IND
Bi, Bismuth (ppm) | 3.7965 | 0.4264 | 3.3381 | 4.2549 | IND | IND
Ca, Calcium (wt.%) | 0.89604 | 0.033182 | 0.87503 | 0.91705 | 0.87508 | 0.917
Ce, Cerium (wt.%) | 1.5274 | 0.052801 | 1.4957 | 1.5591 | 1.489 | 1.5657
Co, Cobalt (ppm) | 18.0159 | 3.3864 | 14.7641 | 21.2677 | 15.5813 | 20.4505
Cr, Chromium (ppm) | 395.1231 | 22.0561 | 381.4415 | 408.8047 | 381.3629 | 408.8833
Cs, Cesium (ppm) | 0.27145 | 0.043588 | 0.23835 | 0.30456 | 0.20715 | 0.33575
Dy, Dysprosium (ppm) | 178.088 | 8.0379 | 173.3844 | 182.7917 | 172.8897 | 183.2864
Er, Erbium (ppm) | 38.1522 | 1.3033 | 37.3983 | 38.9061 | 36.7521 | 39.5524
Eu, Europium (ppm) | 323.8264 | 8.3508 | 320.0393 | 327.6136 | 314.4479 | 333.2049
Fe, Iron (wt.%) | 37.2376 | 1.0071 | 36.4732 | 38.0021 | 36.3759 | 38.0994
Ga, Gallium (ppm) | 105.8017 | 21.1532 | 79.8967 | 131.7066 | 100.7282 | 110.8751
Gd, Gadolinium (ppm) | 676.0411 | 30.2065 | 657.4158 | 694.6664 | 660.9876 | 691.0946
Hf, Hafnium (ppm) | 12.5167 | 0.77046 | 12.0688 | 12.9646 | 11.8488 | 13.1845
Ho, Holmium (ppm) | 21.3049 | 1.3449 | 20.5218 | 22.088 | 20.6623 | 21.9476
In, Indium (ppm) | 2.1418 | 0.16343 | 2.0031 | 2.2804 | IND | IND
K, Potassium (wt.%) | 0.086243 | 0.015822 | 0.069005 | 0.10348 | IND | IND
La, Lanthanum (wt.%) | 1.1706 | 0.031953 | 1.1571 | 1.1842 | 1.142 | 1.1993
Lu, Lutetium (ppm) | 1.6859 | 0.098451 | 1.6402 | 1.7317 | 1.5727 | 1.7992
Mg, Magnesium (wt.%) | 0.84555 | 0.03791 | 0.81806 | 0.87305 | 0.833 | 0.85811
Mn, Manganese (wt.%) | 0.28243 | 0.019215 | 0.26734 | 0.29753 | 0.2724 | 0.29246
Mo, Molybdenum (ppm) | 77.0732 | 4.2684 | 72.7875 | 81.3588 | 73.5556 | 80.5907
Nb, Niobium (ppm) | 1903.4074 | 80.2733 | 1843.1717 | 1963.6431 | 1826.04 | 1980.7749
Nd, Neodymium (wt.%) | 0.99395 | 0.031632 | 0.97757 | 1.0103 | 0.97209 | 1.0158
Ni, Nickel (ppm) | 81.9869 | 15.7398 | 67.4593 | 96.5145 | 66.8339 | 97.1399
P, Phosphorus (wt.%) | 1.4082 | 0.045135 | 1.3724 | 1.444 | 1.3743 | 1.4421
Pb, Lead (ppm) | 173.5236 | 12.3412 | 159.5111 | 187.5361 | 166.6116 | 180.4356
Pr, Praseodymium (ppm) | 2597.1241 | 106.0188 | 2535.3245 | 2658.9237 | 2548.1469 | 2646.1013
Rb, Rubidium (ppm) | 4.3867 | 0.46025 | 4.1715 | 4.6019 | 4.0929 | 4.6804
S, Sulphur (ppm) | 1028.3333 | 90.7092 | 940.048 | 1116.6187 | IND | IND
Sb, Antimony (ppm) | 1.8015 | 0.24593 | 1.4906 | 2.1123 | IND | IND
Si, Silicon (wt.%) | 9.5193 | 0.29384 | 9.2062 | 9.8325 | 9.3261 | 9.7126
Sm, Samarium (ppm) | 1498.3626 | 37.7065 | 1482.1365 | 1514.5888 | 1446.4794 | 1550.2458
Sn, Tin (ppm) | 45.8667 | 2.6766 | 44.329 | 47.4043 | 43.9473 | 47.7861
Sr, Strontium (ppm) | 2060.41 | 69.1644 | 2018.7727 | 2102.0473 | 2016.2531 | 2104.5669
Ta, Tantalum (ppm) | 24.3182 | 3.4672 | 22.0414 | 26.595 | 21.8112 | 26.8252
Tb, Terbium (ppm) | 53.7883 | 2.6975 | 52.2024 | 55.3742 | 52.2359 | 55.3407
Th, Thorium (ppm) | 527.4131 | 23.5659 | 514.0682 | 540.7581 | 513.2255 | 541.6008
Ti, Titanium (wt.%) | 1.9539 | 0.059427 | 1.9089 | 1.9989 | 1.9099 | 1.9978
Tm, Thulium (ppm) | 3.5568 | 0.17456 | 3.4561 | 3.6575 | 3.4396 | 3.674
U, Uranium (ppm) | 17.5665 | 0.41762 | 17.3533 | 17.7798 | 16.9937 | 18.1394
V, Vanadium (ppm) | 230.9526 | 19.0516 | 219.9923 | 241.9129 | 221.5955 | 240.3097
W, Tungsten (ppm) | < 6 | IND | IND | IND | IND | IND
Y, Yttrium (ppm) | 448.5525 | 21.2832 | 435.4316 | 461.6735 | 433.0614 | 464.0437
Yb, Ytterbium (ppm) | 15.7363 | 0.9935 | 15.2593 | 16.2133 | 15.0816 | 16.391
Zn, Zinc (ppm) | 1231.6406 | 95.5595 | 1105.7566 | 1357.5245 | 1185.9131 | 1277.3681
Zr, Zirconium (ppm) | 461.7292 | 17.3791 | 448.099 | 475.3594 | 433.9186 | 489.5397
4-Acid Digestion
Ag, Silver (ppm) | 4.431 | 0.36194 | 4.1744 | 4.6876 | 4.0907 | 4.7713
Al, Aluminium (wt.%) | 4.47 | 0.096259 | 4.4168 | 4.5232 | 4.3795 | 4.5605
As, Arsenic (ppm) | 32.9362 | 4.7153 | 29.0295 | 36.8428 | 31.0003 | 34.8721
Ba, Barium (ppm) | 1648.1311 | 71.5983 | 1609.1976 | 1687.0646 | 1617.2787 | 1678.9835
Be, Beryllium (ppm) | 15.2434 | 1.1937 | 14.4411 | 16.0457 | 14.7762 | 15.7106
Bi, Bismuth (ppm) | 4.0538 | 0.23612 | 3.8951 | 4.2126 | 3.9379 | 4.1697
Ca, Calcium (wt.%) | 0.85095 | 0.058627 | 0.81121 | 0.89068 | 0.83499 | 0.8669
Cd, Cadmium (ppm) | 0.51813 | 0.087485 | 0.4576 | 0.57866 | 0.47175 | 0.56451
Ce, Cerium (wt.%) | 1.5483 | 0.048388 | 1.5011 | 1.5955 | 1.5037 | 1.5929
Co, Cobalt (ppm) | 16.9903 | 0.93218 | 16.6155 | 17.3651 | 16.3818 | 17.5987
Cr, Chromium (ppm) | 334.6808 | 37.2057 | 302.6634 | 366.6982 | 318.7968 | 350.5647
Cs, Cesium (ppm) | 0.28775 | 0.017873 | 0.27526 | 0.30024 | 0.26636 | 0.30913
Cu, Copper (ppm) | 91.8599 | 2.1933 | 90.4906 | 93.2291 | 89.4639 | 94.2559
Dy, Dysprosium (ppm) | 181.7073 | 6.9128 | 177.336 | 186.0787 | 178.5042 | 184.9104
Er, Erbium (ppm) | 36.9362 | 1.5538 | 35.8624 | 38.01 | 36.0763 | 37.796
Eu, Europium (ppm) | 334.4206 | 18.9438 | 321.6848 | 347.1565 | 328.8897 | 339.9516
Fe, Iron (wt.%) | 35.551 | 1.5691 | 34.6214 | 36.4805 | 35.0078 | 36.0942
Gd, Gadolinium (ppm) | 670.037 | 26.3591 | 651.0753 | 688.9987 | 658.0175 | 682.0565
Hf, Hafnium (ppm) | 7.7917 | 1.3232 | 6.7379 | 8.8454 | 7.3504 | 8.2329
Ho, Holmium (ppm) | 21.3127 | 0.93131 | 20.6795 | 21.9459 | 20.8338 | 21.7916
In, Indium (ppm) | 2.1115 | 0.11819 | 2.0351 | 2.1879 | 2.0638 | 2.1592
K, Potassium (wt.%) | 0.081232 | 0.013894 | 0.072375 | 0.090089 | IND | IND
La, Lanthanum (wt.%) | 1.0644 | 0.21298 | 0.85972 | 1.2691 | 1.0328 | 1.096
Li, Lithium (ppm) | 8.2553 | 1.0667 | 7.577 | 8.9337 | 7.9074 | 8.6032
Lu, Lutetium (ppm) | 1.5722 | 0.30173 | 1.3782 | 1.7662 | 1.5095 | 1.6349
Mg, Magnesium (wt.%) | 0.82488 | 0.044558 | 0.79525 | 0.8545 | 0.80877 | 0.84098
Mn, Manganese (wt.%) | 0.26998 | 0.016316 | 0.25843 | 0.28153 | 0.26338 | 0.27658
Mo, Molybdenum (ppm) | 77.135 | 4.3974 | 74.3851 | 79.8848 | 75.4792 | 78.7908
Na, Sodium (wt.%) | 0.14862 | 0.011378 | 0.1419 | 0.15534 | IND | IND
Nd, Neodymium (wt.%) | 0.95892 | 0.025597 | 0.94212 | 0.97572 | 0.94089 | 0.97696
Ni, Nickel (ppm) | 88.5748 | 6.5484 | 84.0044 | 93.1451 | 86.0676 | 91.0819
P, Phosphorus (wt.%) | 1.3595 | 0.066673 | 1.3057 | 1.4133 | 1.3221 | 1.3969
Pb, Lead (ppm) | 204.9508 | 31.7436 | 184.2505 | 225.651 | 199.4657 | 210.4358
Pr, Praseodymium (ppm) | 2600.6204 | 86.8487 | 2515.9774 | 2685.2635 | 2523.7216 | 2677.5193
Rb, Rubidium (ppm) | 4.2208 | 0.39258 | 3.9695 | 4.4722 | 4.0175 | 4.4241
Re, Rhenium (ppm) | < 0.03 | IND | IND | IND | IND | IND
S, Sulphur (ppm) | 983.4111 | 78.8419 | 933.5973 | 1033.2249 | IND | IND
Sb, Antimony (ppm) | 1.868 | 0.1112 | 1.7986 | 1.9375 | 1.7459 | 1.9901
Sc, Scandium (ppm) | 141.3799 | 5.8076 | 137.418 | 145.3418 | 138.1328 | 144.6271
Se, Selenium (ppm) | < 50 | IND | IND | IND | IND | IND
Sm, Samarium (ppm) | 1484.8664 | 49.3775 | 1437.6604 | 1532.0725 | 1437.1586 | 1532.5743
Sn, Tin (ppm) | 35.6779 | 3.2051 | 33.5317 | 37.824 | 34.291 | 37.0647
Sr, Strontium (ppm) | 2042.8066 | 144.2724 | 1943.8925 | 2141.7206 | 2010.4634 | 2075.1497
Ta, Tantalum (ppm) | 20.9121 | 2.6948 | 18.6003 | 23.2239 | 19.7803 | 22.0439
Tb, Terbium (ppm) | 54.8891 | 2.5409 | 53.0687 | 56.7095 | 53.4438 | 56.3343
Te, Tellurium (ppm) | 0.3625 | 0.067883 | 0.30112 | 0.42387 | 0.32874 | 0.39625
Th, Thorium (ppm) | 506.3839 | 59.517 | 458.8648 | 553.903 | 496.0409 | 516.727
Ti, Titanium (wt.%) | < 1 | IND | IND | IND | IND | IND
Tl, Thallium (ppm) | 0.14535 | 0.012814 | 0.13668 | 0.15402 | IND | IND
Tm, Thulium (ppm) | 3.2956 | 0.3212 | 3.0776 | 3.5137 | 3.1935 | 3.3978
U, Uranium (ppm) | 17.6356 | 0.61435 | 17.2368 | 18.0343 | 17.1632 | 18.1079
V, Vanadium (ppm) | 207.0274 | 26.1739 | 187.7429 | 226.3118 | 202.5423 | 211.5124
W, Tungsten (ppm) | 2.3747 | 0.40124 | 2.1047 | 2.6447 | 2.2443 | 2.505
Y, Yttrium (ppm) | 448.8679 | 33.2729 | 424.5828 | 473.153 | 442.5434 | 455.1925
Yb, Ytterbium (ppm) | 14.277 | 0.93752 | 13.6036 | 14.9503 | 13.7797 | 14.7742
Zn, Zinc (ppm) | 1128.0782 | 56.7693 | 1090.3079 | 1165.8486 | 1110.3645 | 1145.792
Zr, Zirconium (ppm) | 209.7096 | 39.9025 | 179.0074 | 240.4119 | 200.3444 | 219.0749

## Indicative Values
Table 3. Indicative Values for OREAS 464
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Al | wt.% | 4.6922 | Lu | ppm | < 20 | Sr | ppm | 2001.2402
Ba | ppm | 1813.7035 | Mg | wt.% | 0.85509 | Ta | ppm | < 100
Ca | wt.% | 0.91511 | Mn | wt.% | 0.30592 | Tb | ppm | 49.4229
Cr | ppm | 325.5787 | Na | wt.% | 0.17396 | Ti | wt.% | 1.9706
Dy | ppm | 172.0833 | Nb | ppm | 1938.2384 | Tm | ppm | < 10
Er | ppm | 34.9803 | Ni | ppm | 157.1709 | U | ppm | < 100
Eu | ppm | 287.8775 | P | wt.% | 1.4202 | V | ppm | 257.6742
Gd | ppm | 697.698 | Pb | ppm | 278.4998 | W | ppm | < 100
Hf | ppm | < 100 | Si | wt.% | 9.354 | Yb | ppm | 17.5639
Ho | ppm | 17.4596 | Sn | ppm | < 100 | Zn | ppm | 1258.569
K | wt.% | 0.085436 | S | wt.% | 0.099378 | Zr | ppm | 456.5196
Thermogravimetry
H2O- | wt.% | 0.50333
Borate / Peroxide Fusion ICP
Ag | ppm | 10.4167 | Ge | ppm | 43.3333 | Se | ppm | 40
As | ppm | 251.8708 | Li | ppm | 8.1337 | Te | ppm | < 1
B | ppm | 107.5833 | Na | wt.% | 0.13501 | Tl | ppm | < 0.5
Cd | ppm | < 1 | Re | ppm | < 0.1
Cu | ppm | 103.6968 | Sc | ppm | 151.875
4-Acid Digestion
Ga | ppm | 97.0371 | Ge | ppm | 18.3781 | Nb | ppm | 1597.7592

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 464
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Ce, wt.% | 1.5426 | 0.022149 | 1.4983 | 1.5869 | 1.4762 | 1.6091 | 1.436% | 2.872% | 4.307% | 1.4655 | 1.6198
Fe, wt.% | 38.0367 | 0.72116 | 36.5943 | 39.479 | 35.8732 | 40.2001 | 1.896% | 3.792% | 5.688% | 36.1348 | 39.9385
La, wt.% | 1.1991 | 0.01683 | 1.1655 | 1.2328 | 1.1487 | 1.2496 | 1.404% | 2.807% | 4.211% | 1.1392 | 1.2591
Nd, wt.% | 0.96367 | 0.030982 | 0.90171 | 1.0256 | 0.87073 | 1.0566 | 3.215% | 6.43% | 9.645% | 0.91549 | 1.0119
Pr, ppm | 2540.6949 | 81.564 | 2377.5669 | 2703.8229 | 2296.0029 | 2785.3869 | 3.21% | 6.421% | 9.631% | 2413.6602 | 2667.7296
Sm, ppm | 1409.7243 | 47.6299 | 1314.4645 | 1504.9841 | 1266.8346 | 1552.614 | 3.379% | 6.757% | 10.14% | 1339.2381 | 1480.2105
Th, ppm | 506.0026 | 74.2303 | 357.5421 | 654.4632 | 283.3118 | 728.6935 | 14.67% | 29.34% | 44.01% | 480.7025 | 531.3028
Y, ppm | 442.4618 | 54.1764 | 334.1089 | 550.8147 | 279.9324 | 604.9911 | 12.24% | 24.49% | 36.73% | 420.3387 | 464.5849
Thermogravimetry
LOI, wt.% | 0.6969 | 0.0711 | 0.5547 | 0.8391 | 0.4836 | 0.9102 | 10.2% | 20.4% | 30.61% | 0.66205 | 0.73174
Borate / Peroxide Fusion ICP
Al, wt.% | 4.6305 | 0.15029 | 4.3299 | 4.931 | 4.1796 | 5.0813 | 3.246% | 6.491% | 9.737% | 4.3989 | 4.862
Ba, ppm | 1656.7244 | 69.2874 | 1518.1495 | 1795.2992 | 1448.8621 | 1864.5866 | 4.182% | 8.364% | 12.55% | 1573.8882 | 1739.5606
Be, ppm | 15.6289 | 1.5763 | 12.4763 | 18.7815 | 10.9 | 20.3578 | 10.09% | 20.17% | 30.26% | 14.8474 | 16.4103
Bi, ppm | 3.7965 | 0.4264 | 2.9437 | 4.6493 | 2.5173 | 5.0757 | 11.23% | 22.46% | 33.69% | 3.6067 | 3.9863
Ca, wt.% | 0.89604 | 0.033182 | 0.82967 | 0.9624 | 0.79649 | 0.99558 | 3.703% | 7.406% | 11.11% | 0.85124 | 0.94084
Ce, wt.% | 1.5274 | 0.052801 | 1.4218 | 1.633 | 1.369 | 1.6858 | 3.457% | 6.914% | 10.37% | 1.451 | 1.6038
Co, ppm | 18.0159 | 3.3864 | 11.2431 | 24.7887 | 7.8567 | 28.1751 | 18.8% | 37.59% | 56.39% | 17.1151 | 18.9167
Cr, ppm | 395.1231 | 22.0561 | 351.0109 | 439.2353 | 328.9548 | 461.2914 | 5.582% | 11.16% | 16.75% | 375.3669 | 414.8792
Cs, ppm | 0.27145 | 0.043588 | 0.18428 | 0.35863 | 0.14069 | 0.40222 | 16.06% | 32.11% | 48.17% | 0.25788 | 0.28502
Dy, ppm | 178.088 | 8.0379 | 162.0123 | 194.1638 | 153.9744 | 202.2017 | 4.513% | 9.027% | 13.54% | 169.1836 | 186.9924
Er, ppm | 38.1522 | 1.3033 | 35.5455 | 40.7589 | 34.2422 | 42.0623 | 3.416% | 6.832% | 10.25% | 36.2446 | 40.0598
Eu, ppm | 323.8264 | 8.3508 | 307.1248 | 340.5281 | 298.774 | 348.8789 | 2.579% | 5.158% | 7.736% | 307.6351 | 340.0178
Fe, wt.% | 37.2376 | 1.0071 | 35.2235 | 39.2518 | 34.2164 | 40.2589 | 2.704% | 5.409% | 8.113% | 35.3758 | 39.0995
Ga, ppm | 105.8017 | 21.1532 | 63.4953 | 148.108 | 42.3421 | 169.2612 | 19.99% | 39.99% | 59.98% | 100.5116 | 111.0917
Gd, ppm | 676.0411 | 30.2065 | 615.628 | 736.4542 | 585.4215 | 766.6607 | 4.468% | 8.936% | 13.4% | 642.239 | 709.8432
Hf, ppm | 12.5167 | 0.77046 | 10.9757 | 14.0576 | 10.2053 | 14.8281 | 6.156% | 12.31% | 18.47% | 11.8908 | 13.1425
Ho, ppm | 21.3049 | 1.3449 | 18.615 | 23.9948 | 17.2701 | 25.3397 | 6.313% | 12.63% | 18.94% | 20.2397 | 22.3702
In, ppm | 2.1418 | 0.16343 | 1.8149 | 2.4686 | 1.6515 | 2.632 | 7.63% | 15.26% | 22.89% | 2.0347 | 2.2489
K, wt.% | 0.086243 | 0.015822 | 0.054598 | 0.11789 | 0.038776 | 0.13371 | 18.35% | 36.69% | 55.04% | 0.081931 | 0.090555
La, wt.% | 1.1706 | 0.031953 | 1.1067 | 1.2345 | 1.0748 | 1.2665 | 2.73% | 5.459% | 8.189% | 1.1121 | 1.2291
Lu, ppm | 1.6859 | 0.098451 | 1.489 | 1.8829 | 1.3906 | 1.9813 | 5.839% | 11.68% | 17.52% | 1.6017 | 1.7702
Mg, wt.% | 0.84555 | 0.03791 | 0.76973 | 0.92137 | 0.73182 | 0.95928 | 4.483% | 8.967% | 13.45% | 0.80327 | 0.88783
Mn, wt.% | 0.28243 | 0.019215 | 0.244 | 0.32086 | 0.22479 | 0.34008 | 6.804% | 13.61% | 20.41% | 0.26831 | 0.29655
Mo, ppm | 77.0732 | 4.2684 | 68.5365 | 85.6099 | 64.2681 | 89.8782 | 5.538% | 11.08% | 16.61% | 73.2195 | 80.9268
Nb, ppm | 1903.4074 | 80.2733 | 1742.8608 | 2063.954 | 1662.5875 | 2144.2273 | 4.217% | 8.435% | 12.65% | 1808.237 | 1998.5778
Nd, wt.% | 0.99395 | 0.031632 | 0.93069 | 1.0572 | 0.89906 | 1.0889 | 3.182% | 6.365% | 9.547% | 0.94426 | 1.0437
Ni, ppm | 81.9869 | 15.7398 | 50.5072 | 113.4665 | 34.7674 | 129.2064 | 19.2% | 38.4% | 57.59% | 77.8876 | 86.0862
P, wt.% | 1.4082 | 0.045135 | 1.318 | 1.4985 | 1.2728 | 1.5436 | 3.205% | 6.41% | 9.615% | 1.3378 | 1.4786
Pb, ppm | 173.5236 | 12.3412 | 148.8413 | 198.2059 | 136.5002 | 210.5471 | 7.112% | 14.22% | 21.34% | 164.8474 | 182.1998
Pr, ppm | 2597.1241 | 106.0188 | 2385.0865 | 2809.1617 | 2279.0677 | 2915.1805 | 4.082% | 8.164% | 12.25% | 2467.2679 | 2726.9803
Rb, ppm | 4.3867 | 0.46025 | 3.4662 | 5.3072 | 3.0059 | 5.7674 | 10.49% | 20.98% | 31.48% | 4.1673 | 4.606
S, ppm | 1028.3333 | 90.7092 | 846.9149 | 1209.7518 | 756.2057 | 1300.461 | 8.821% | 17.64% | 26.46% | 976.9167 | 1079.75
Sb, ppm | 1.8015 | 0.24593 | 1.3096 | 2.2933 | 1.0637 | 2.5392 | 13.65% | 27.3% | 40.95% | 1.7114 | 1.8915
Si, wt.% | 9.5193 | 0.29384 | 8.9317 | 10.107 | 8.6378 | 10.4009 | 3.087% | 6.174% | 9.26% | 9.0434 | 9.9953
Sm, ppm | 1498.3626 | 37.7065 | 1422.9495 | 1573.7757 | 1385.243 | 1611.4823 | 2.517% | 5.033% | 7.55% | 1423.4445 | 1573.2808
Sn, ppm | 45.8667 | 2.6766 | 40.5135 | 51.2198 | 37.837 | 53.8964 | 5.836% | 11.67% | 17.51% | 43.5733 | 48.16
Sr, ppm | 2060.41 | 69.1644 | 1922.0813 | 2198.7387 | 1852.9169 | 2267.9031 | 3.357% | 6.714% | 10.07% | 1957.3895 | 2163.4305
Ta, ppm | 24.3182 | 3.4672 | 17.3838 | 31.2526 | 13.9166 | 34.7198 | 14.26% | 28.52% | 42.77% | 23.1023 | 25.5341
Tb, ppm | 53.7883 | 2.6975 | 48.3932 | 59.1834 | 45.6957 | 61.8809 | 5.015% | 10.03% | 15.05% | 51.0989 | 56.4777
Th, ppm | 527.4131 | 23.5659 | 480.2814 | 574.5448 | 456.7156 | 598.1107 | 4.468% | 8.936% | 13.4% | 501.0425 | 553.7838
Ti, wt.% | 1.9539 | 0.059427 | 1.835 | 2.0727 | 1.7756 | 2.1322 | 3.041% | 6.083% | 9.124% | 1.8562 | 2.0516
Tm, ppm | 3.5568 | 0.17456 | 3.2077 | 3.9059 | 3.0331 | 4.0805 | 4.908% | 9.815% | 14.72% | 3.379 | 3.7346
U, ppm | 17.5665 | 0.41762 | 16.7313 | 18.4018 | 16.3137 | 18.8194 | 2.377% | 4.755% | 7.132% | 16.6882 | 18.4449
V, ppm | 230.9526 | 19.0516 | 192.8495 | 269.0557 | 173.798 | 288.1073 | 8.249% | 16.5% | 24.75% | 219.405 | 242.5003
W, ppm | < 6 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Y, ppm | 448.5525 | 21.2832 | 405.9861 | 491.119 | 384.7028 | 512.4023 | 4.745% | 9.49% | 14.23% | 426.1249 | 470.9802
Yb, ppm | 15.7363 | 0.9935 | 13.7493 | 17.7233 | 12.7558 | 18.7168 | 6.313% | 12.63% | 18.94% | 14.9495 | 16.5231
Zn, ppm | 1231.6406 | 95.5595 | 1040.5216 | 1422.7595 | 944.9622 | 1518.319 | 7.759% | 15.52% | 23.28% | 1170.0586 | 1293.2226
Zr, ppm | 461.7292 | 17.3791 | 426.9711 | 496.4873 | 409.592 | 513.8663 | 3.764% | 7.528% | 11.29% | 438.6427 | 484.8156
4-Acid Digestion
Ag, ppm | 4.431 | 0.36194 | 3.7071 | 5.1549 | 3.3452 | 5.5168 | 8.168% | 16.34% | 24.51% | 4.2094 | 4.6525
Al, wt.% | 4.47 | 0.096259 | 4.2775 | 4.6625 | 4.1812 | 4.7588 | 2.153% | 4.307% | 6.46% | 4.2465 | 4.6935
As, ppm | 32.9362 | 4.7153 | 23.5056 | 42.3668 | 18.7903 | 47.0821 | 14.32% | 28.63% | 42.95% | 31.2894 | 34.583
Ba, ppm | 1648.1311 | 71.5983 | 1504.9345 | 1791.3277 | 1433.3362 | 1862.9261 | 4.344% | 8.688% | 13.03% | 1565.7246 | 1730.5377
Be, ppm | 15.2434 | 1.1937 | 12.8559 | 17.6309 | 11.6622 | 18.8246 | 7.831% | 15.66% | 23.49% | 14.4812 | 16.0056
Bi, ppm | 4.0538 | 0.23612 | 3.5816 | 4.5261 | 3.3455 | 4.7622 | 5.825% | 11.65% | 17.47% | 3.8511 | 4.2565
Ca, wt.% | 0.85095 | 0.058627 | 0.73369 | 0.9682 | 0.67507 | 1.0268 | 6.89% | 13.78% | 20.67% | 0.8084 | 0.89349
Cd, ppm | 0.51813 | 0.087485 | 0.34316 | 0.6931 | 0.25567 | 0.78059 | 16.88% | 33.77% | 50.65% | 0.49222 | 0.54404
Ce, wt.% | 1.5483 | 0.048388 | 1.4515 | 1.6451 | 1.4031 | 1.6934 | 3.125% | 6.251% | 9.376% | 1.4709 | 1.6257
Co, ppm | 16.9903 | 0.93218 | 15.1259 | 18.8546 | 14.1937 | 19.7868 | 5.487% | 10.97% | 16.46% | 16.1408 | 17.8398
Cr, ppm | 334.6808 | 37.2057 | 260.2693 | 409.0923 | 223.0636 | 446.298 | 11.12% | 22.23% | 33.35% | 317.9468 | 351.4148
Cs, ppm | 0.28775 | 0.017873 | 0.252 | 0.3235 | 0.23413 | 0.34137 | 6.211% | 12.42% | 18.63% | 0.27336 | 0.30214
Cu, ppm | 91.8599 | 2.1933 | 87.4734 | 96.2464 | 85.2801 | 98.4397 | 2.388% | 4.775% | 7.163% | 87.2669 | 96.4529
Dy, ppm | 181.7073 | 6.9128 | 167.8817 | 195.533 | 160.9688 | 202.4458 | 3.804% | 7.609% | 11.41% | 172.622 | 190.7927
Er, ppm | 36.9362 | 1.5538 | 33.8286 | 40.0437 | 32.2749 | 41.5975 | 4.207% | 8.413% | 12.62% | 35.0894 | 38.783
Eu, ppm | 334.4206 | 18.9438 | 296.5331 | 372.3082 | 277.5893 | 391.2519 | 5.665% | 11.33% | 16.99% | 317.6996 | 351.1417
Fe, wt.% | 35.551 | 1.5691 | 32.4127 | 38.6892 | 30.8436 | 40.2583 | 4.414% | 8.827% | 13.24% | 33.7734 | 37.3285
Gd, ppm | 670.037 | 26.3591 | 617.3189 | 722.7551 | 590.9598 | 749.1142 | 3.934% | 7.868% | 11.8% | 636.5352 | 703.5389
Hf, ppm | 7.7917 | 1.3232 | 5.1454 | 10.438 | 3.8222 | 11.7611 | 16.98% | 33.96% | 50.94% | 7.4021 | 8.1813
Ho, ppm | 21.3127 | 0.93131 | 19.4501 | 23.1753 | 18.5188 | 24.1066 | 4.37% | 8.739% | 13.11% | 20.2471 | 22.3783
In, ppm | 2.1115 | 0.11819 | 1.8751 | 2.3479 | 1.757 | 2.4661 | 5.597% | 11.19% | 16.79% | 2.0059 | 2.2171
K, wt.% | 0.081232 | 0.013894 | 0.053444 | 0.10902 | 0.039551 | 0.12291 | 17.1% | 34.21% | 51.31% | 0.07717 | 0.085294
La, wt.% | 1.0644 | 0.21298 | 0.63846 | 1.4904 | 0.42548 | 1.7034 | 20.01% | 40.02% | 60.03% | 1.0112 | 1.1176
Li, ppm | 8.2553 | 1.0667 | 6.122 | 10.3887 | 5.0553 | 11.4553 | 12.92% | 25.84% | 38.76% | 7.8426 | 8.6681
Lu, ppm | 1.5722 | 0.30173 | 0.9687 | 2.1756 | 0.66696 | 2.4774 | 19.19% | 38.38% | 57.58% | 1.4936 | 1.6508
Mg, wt.% | 0.82488 | 0.044558 | 0.73576 | 0.91399 | 0.69121 | 0.95855 | 5.402% | 10.8% | 16.21% | 0.78363 | 0.86612
Mn, wt.% | 0.26998 | 0.016316 | 0.23735 | 0.30261 | 0.22103 | 0.31893 | 6.043% | 12.09% | 18.13% | 0.25648 | 0.28348
Mo, ppm | 77.135 | 4.3974 | 68.3401 | 85.9298 | 63.9427 | 90.3273 | 5.701% | 11.4% | 17.1% | 73.2782 | 80.9917
Na, wt.% | 0.14862 | 0.011378 | 0.12587 | 0.17138 | 0.11449 | 0.18275 | 7.655% | 15.31% | 22.97% | 0.14119 | 0.15605
Nd, wt.% | 0.95892 | 0.025597 | 0.90773 | 1.0101 | 0.88213 | 1.0357 | 2.669% | 5.339% | 8.008% | 0.91098 | 1.0069
Ni, ppm | 88.5748 | 6.5484 | 75.4779 | 101.6716 | 68.9294 | 108.2201 | 7.393% | 14.79% | 22.18% | 84.146 | 93.0035
P, wt.% | 1.3595 | 0.066673 | 1.2262 | 1.4929 | 1.1595 | 1.5595 | 4.904% | 9.808% | 14.71% | 1.2916 | 1.4275
Pb, ppm | 204.9508 | 31.7436 | 141.4637 | 268.4379 | 109.7201 | 300.1814 | 15.49% | 30.98% | 46.47% | 194.7032 | 215.1983
Pr, ppm | 2600.6204 | 86.8487 | 2426.923 | 2774.3179 | 2340.0742 | 2861.1666 | 3.34% | 6.679% | 10.02% | 2470.5894 | 2730.6515
Rb, ppm | 4.2208 | 0.39258 | 3.4357 | 5.006 | 3.0431 | 5.3986 | 9.301% | 18.6% | 27.9% | 4.0098 | 4.4319
Re, ppm | < 0.03 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
S, ppm | 983.4111 | 78.8419 | 825.7273 | 1141.0949 | 746.8854 | 1219.9369 | 8.017% | 16.03% | 24.05% | 934.2406 | 1032.5817
Sb, ppm | 1.868 | 0.1112 | 1.6456 | 2.0904 | 1.5344 | 2.2016 | 5.953% | 11.91% | 17.86% | 1.7746 | 1.9614
Sc, ppm | 141.3799 | 5.8076 | 129.7648 | 152.9951 | 123.9572 | 158.8026 | 4.108% | 8.216% | 12.32% | 134.3109 | 148.4489
Se, ppm | < 50 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sm, ppm | 1484.8664 | 49.3775 | 1386.1114 | 1583.6215 | 1336.7339 | 1632.999 | 3.325% | 6.651% | 9.976% | 1410.6231 | 1559.1097
Sn, ppm | 35.6779 | 3.2051 | 29.2676 | 42.0881 | 26.0625 | 45.2932 | 8.983% | 17.97% | 26.95% | 33.894 | 37.4617
Sr, ppm | 2042.8066 | 144.2724 | 1754.2618 | 2331.3513 | 1609.9894 | 2475.6237 | 7.062% | 14.12% | 21.19% | 1940.6662 | 2144.9469
Ta, ppm | 20.9121 | 2.6948 | 15.5225 | 26.3017 | 12.8278 | 28.9965 | 12.89% | 25.77% | 38.66% | 19.8665 | 21.9577
Tb, ppm | 54.8891 | 2.5409 | 49.8073 | 59.9708 | 47.2664 | 62.5117 | 4.629% | 9.258% | 13.89% | 52.1446 | 57.6335
Te, ppm | 0.3625 | 0.067883 | 0.22673 | 0.49826 | 0.15885 | 0.56615 | 18.73% | 37.45% | 56.18% | 0.34437 | 0.38062
Th, ppm | 506.3839 | 59.517 | 387.35 | 625.4179 | 327.833 | 684.9349 | 11.75% | 23.51% | 35.26% | 481.0647 | 531.7031
Ti, wt.% | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Tl, ppm | 0.14535 | 0.012814 | 0.11972 | 0.17098 | 0.10691 | 0.18379 | 8.816% | 17.63% | 26.45% | 0.13808 | 0.15262
Tm, ppm | 3.2956 | 0.3212 | 2.6532 | 3.938 | 2.332 | 4.2592 | 9.746% | 19.49% | 29.24% | 3.1309 | 3.4604
U, ppm | 17.6356 | 0.61435 | 16.4069 | 18.8642 | 15.7925 | 19.4786 | 3.484% | 6.967% | 10.45% | 16.7538 | 18.5173
V, ppm | 207.0274 | 26.1739 | 154.6796 | 259.3751 | 128.5057 | 285.549 | 12.64% | 25.29% | 37.93% | 196.676 | 217.3787
W, ppm | 2.3747 | 0.40124 | 1.5722 | 3.1772 | 1.171 | 3.5784 | 16.9% | 33.79% | 50.69% | 2.256 | 2.4934
Y, ppm | 448.8679 | 33.2729 | 382.3221 | 515.4138 | 349.0491 | 548.6867 | 7.413% | 14.83% | 22.24% | 426.4245 | 471.3113
Yb, ppm | 14.277 | 0.93752 | 12.4019 | 16.152 | 11.4644 | 17.0895 | 6.567% | 13.13% | 19.7% | 13.5631 | 14.9908
Zn, ppm | 1128.0782 | 56.7693 | 1014.5396 | 1241.6169 | 957.7703 | 1298.3862 | 5.032% | 10.06% | 15.1% | 1071.6743 | 1184.4822
Zr, ppm | 209.7096 | 39.9025 | 129.9046 | 289.5146 | 90.0021 | 329.4171 | 19.03% | 38.05% | 57.08% | 199.2242 | 220.1951

## Fusion XRF
| Analytical results for Al in OREAS 464 (Indicative Value 4.69 wt.%) | [8 empty] | Indicative
Al | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 4.726 | 4.615 | 4.689 | 4.747 |  |  |  | 1
| 1 | 2 | 4.716 | 4.589 | 4.673 | 4.8 |  |  |  | 1
| 1 | 3 | 4.678 | 4.631 | 4.678 | 4.774 |  |  |  | 16
| 1 | 4 | 4.678 | 4.62 | 4.71 | 4.779 |  |  |  | 4.6922
| 1 | 5 | 4.726 | 4.589 | 4.742 | 4.742 |  |  |  | 7
| 1 | 6 | 4.694 | 4.562 | 4.694 | 4.758
| Mean |  | 4.703 | 4.601 | 4.6977 | 4.7667
| Median |  | 4.705 | 4.602 | 4.6915 | 4.766
| Std Dev. |  | 0.022619 | 0.025558 | 0.025303 | 0.021851
| Rel.Std.Dev. |  | 0.4809% | 0.5555% | 0.5386% | 0.4584%
| PDM3 |  | 0.2311% | -1.943% | 0.1174% | 1.588%
| Z-Score (Absolute) |  | 0.05 | 1.95 | 0.05 | 1.3
| Analytical results for Ba in OREAS 464 (Indicative Value 1814 ppm) | [8 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 1790 | 1790 | [5 empty] | 1
| 1 | 2 | 1970 | 1790 | [5 empty] | 2
| 1 | 3 | 1880 | 1790 | [5 empty] | 16
| 1 | 4 | 1880 | 1790 | [5 empty] | 1813.7035
| 1 | 5 | 1790 | 1700 | [5 empty] | 8
| 1 | 6 | 1700 | 1880
| Mean |  | 1835 | 1790
| Median |  | 1835 | 1790
| Std Dev. |  | 94.3928 | 56.921
| Rel.Std.Dev. |  | 5.144% | 3.18%
| PDM3 |  | 1.174% | -1.307%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ca in OREAS 464 (Indicative Value 0.915 wt.%) | [8 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.901 | 0.922 | 0.879 | 0.951 |  |  |  | 1
| 1 | 2 | 0.908 | 0.901 | 0.901 | 0.951 |  |  |  | 3
| 1 | 3 | 0.908 | 0.908 | 0.915 | 0.943 |  |  |  | 16
| 1 | 4 | 0.901 | 0.936 | 0.879 | 0.951 |  |  |  | 0.91511
| 1 | 5 | 0.901 | 0.929 | 0.893 | 0.943 |  |  |  | 9
| 1 | 6 | 0.893 | 0.929 | 0.872 | 0.951
| Mean |  | 0.902 | 0.92083 | 0.88983 | 0.94833
| Median |  | 0.901 | 0.9255 | 0.886 | 0.951
| Std Dev. |  | 0.0055857 | 0.013586 | 0.016253 | 0.0041312
| Rel.Std.Dev. |  | 0.6193% | 1.475% | 1.827% | 0.4356%
| PDM3 |  | -1.432% | 0.6258% | -2.762% | 3.631%
| Z-Score (Absolute) |  | 0.41 | 0.41 | 0.93 | 1.61
| Analytical results for Ce in OREAS 464 (Certified Value 1.54 wt.%) | [8 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.563 | 1.563 | 1.528 | 1.522 | 1.631 | 1.55 | 1.514 | 1
| 1 | 2 | 1.579 | 1.563 | 1.511 | 1.547 | 1.631 | 1.55 | 1.498 | 0
| 1 | 3 | 1.571 | 1.563 | 1.511 | 1.547 | 1.622 | 1.55 | 1.506 | 16
| 1 | 4 | 1.579 | 1.547 | 1.537 | 1.53 | 1.622 | 1.55 | 1.506 | 1.5426
| 1 | 5 | 1.555 | 1.563 | 1.571 | 1.539 | 1.597 | 1.55 | 1.522 | 7
| 1 | 6 | 1.563 | 1.555 | 1.537 | 1.53 | 1.622 | 1.55 | 1.514
| Mean |  | 1.5683 | 1.559 | 1.5325 | 1.5358 | 1.6208 | 1.55 | 1.51
| Median |  | 1.567 | 1.563 | 1.5325 | 1.5345 | 1.622 | 1.55 | 1.51
| Std Dev. |  | 0.0096885 | 0.0066933 | 0.022215 | 0.010187 | 0.012481 | 0 | 0.0083905
| Rel.Std.Dev. |  | 0.6178% | 0.4293% | 1.45% | 0.6633% | 0.77% | 0% | 0.5557%
| PDM3 |  | 1.666% | 1.061% | -0.6572% | -0.4411% | 5.069% | 0.4772% | -2.116%
| Z-Score (Absolute) |  | 0.71 | 0.34 | 0.67 | 0.54 | 2.91 | 0 | 1.54
| Analytical results for Cr in OREAS 464 (Indicative Value 326 ppm) | [8 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 480 | 205 | [5 empty] | 1
| 1 | 2 | 410 | 226 | [5 empty] | 4
| 1 | 3 | 410 | 226 | [5 empty] | 16
| 1 | 4 | 410 | 205 | [5 empty] | 325.5787
| 1 | 5 | 480 | 212 | [5 empty] | 10
| 1 | 6 | 410 | 233
| Mean |  | 433.3333 | 217.8333
| Median |  | 410 | 219
| Std Dev. |  | 36.1478 | 12.0568
| Rel.Std.Dev. |  | 8.342% | 5.535%
| PDM3 |  | 33.1% | -33.09%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Dy in OREAS 464 (Indicative Value 172 ppm) | [8 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 260 | 170 | 166 |  |  |  |  | 1
| 1 | 2 | 260 | 170 | 174 |  |  |  |  | 5
| 1 | 3 | 260 | 260 | 166 |  |  |  |  | 16
| 1 | 4 | 260 | 170 | 166 |  |  |  |  | 172.0833
| 1 | 5 | 260 | 170 | 174 |  |  |  |  | 11
| 1 | 6 | 260 | 170 | 174
| Mean |  | 260 | 185 | 170
| Median |  | 260 | 170 | 170
| Std Dev. |  | 0 | 36.7423 | 4.3818
| Rel.Std.Dev. |  | 0% | 19.86% | 2.578%
| PDM3 |  | 51.09% | 7.506% | -1.211%
| Z-Score (Absolute) |  | 13.27 | 0 | 0.67
| Analytical results for Er in OREAS 464 (Indicative Value 35 ppm) | [8 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 90 | < 90 | 26 |  |  |  |  | 1
| 1 | 2 | < 90 | 90 | 26 |  |  |  |  | 6
| 1 | 3 | < 90 | < 90 | 26 |  |  |  |  | 16
| 1 | 4 | < 90 | < 90 | 26 |  |  |  |  | 34.9803
| 1 | 5 | < 90 | < 90 | 26 |  |  |  |  | 12
| 1 | 6 | < 90 | < 90 | 26
| Mean |  |  | 90 | 26
| Median |  |  | 90 | 26
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
| PDM3 |  |  | 157.3% | -25.67%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Eu in OREAS 464 (Indicative Value 288 ppm) | [8 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 09 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 260 | 260 | 260 |  |  |  |  | 1
| 1 | 2 | 350 | 260 | 260 |  |  |  |  | 7
| 1 | 3 | 350 | 260 | 260 |  |  |  |  | 16
| 1 | 4 | 350 | 260 | 260 |  |  |  |  | 287.8775
| 1 | 5 | 350 | 260 | 260 |  |  |  |  | 13
| 1 | 6 | 350 | 260 | 260
| Mean |  | 335 | 260 | 260
| Median |  | 350 | 260 | 260
| Std Dev. |  | 36.7423 | 0 | 0
| Rel.Std.Dev. |  | 10.97% | 0% | 0%
| PDM3 |  | 16.37% | -9.684% | -9.684%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Fe in OREAS 464 (Certified Value 38.04 wt.%) | [8 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 09 | 13 | 14 | 15 | 22 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 38.155 | 38.819 | 37.07 | 37.35 | 38.54 |  |  | 1
| 1 | 2 | 38.134 | 38.819 | 37.28 | 37.281 | 38.68 |  |  | 0
| 1 | 3 | 38.358 | 39.099 | 37.21 | 37.281 | 38.47 |  |  | 16
| 1 | 4 | 38.148 | 39.099 | 37.28 | 37.211 | 38.47 |  |  | 38.0367
| 1 | 5 | 38.288 | 39.029 | 37.42 | 37.141 | 38.68 |  |  | 8
| 1 | 6 | 38.008 | 38.959 | 37.07 | 37.211 | 38.54
| Mean |  | 38.1818 | 38.9707 | 37.2217 | 37.2458 | 38.5633
| Median |  | 38.1515 | 38.994 | 37.245 | 37.246 | 38.54
| Std Dev. |  | 0.12384 | 0.12844 | 0.13586 | 0.073131 | 0.095638
| Rel.Std.Dev. |  | 0.3243% | 0.3296% | 0.365% | 0.1963% | 0.248%
| PDM3 |  | 0.3817% | 2.456% | -2.143% | -2.079% | 1.385%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.82 | 0.8 | 0.33
| Analytical results for Gd in OREAS 464 (Indicative Value 698 ppm) | [8 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 09 | 14 | 22 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 690 | 690 | 610 | 690 |  |  |  | 1
| 1 | 2 | 690 | 690 | 690 | 690 |  |  |  | 8
| 1 | 3 | 690 | 780 | 690 | 690 |  |  |  | 16
| 1 | 4 | 690 | 780 | 610 | 690 |  |  |  | 697.698
| 1 | 5 | 690 | 780 | 690 | 690 |  |  |  | 14
| 1 | 6 | 690 | 690 | 690 | 690
| Mean |  | 690 | 735 | 663.3333 | 690
| Median |  | 690 | 735 | 690 | 690
| Std Dev. |  | 0 | 49.295 | 41.3118 | 0
| Rel.Std.Dev. |  | 0% | 6.707% | 6.228% | 0%
| PDM3 |  | -1.103% | 5.346% | -4.925% | -1.103%
| Z-Score (Absolute) |  | 0 | 2.02 | 1.35 | 0
| Analytical results for Hf in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Hf | Round | Replicate | Lab | [6 empty] | 1
| No | No | 13 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 9
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 15
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 464 (Indicative Value 17.5 ppm) | [8 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | 17 | [5 empty] | 1
| 1 | 2 | < 90 | 17 | [5 empty] | 10
| 1 | 3 | < 90 | 17 | [5 empty] | 16
| 1 | 4 | < 90 | 17 | [5 empty] | 17.4596
| 1 | 5 | < 90 | 17 | [5 empty] | 16
| 1 | 6 | < 90 | 17
| Mean |  |  | 17
| Median |  |  | 17
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | -2.632%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 464 (Indicative Value 0.085 wt.%) | [8 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.083 | 0.1 | 0.075 | 0.083 |  |  |  | 1
| 1 | 2 | 0.083 | 0.108 | 0.075 | 0.091 |  |  |  | 11
| 1 | 3 | 0.083 | 0.091 | 0.083 | 0.083 |  |  |  | 16
| 1 | 4 | 0.083 | 0.1 | 0.075 | 0.083 |  |  |  | 0.085436
| 1 | 5 | 0.083 | 0.1 | 0.075 | 0.083 |  |  |  | 17
| 1 | 6 | 0.083 | 0.1 | 0.075 | 0.083
| Mean |  | 0.083 | 0.099833 | 0.076333 | 0.084333
| Median |  | 0.083 | 0.1 | 0.075 | 0.083
| Std Dev. |  | 0 | 0.0053821 | 0.003266 | 0.003266
| Rel.Std.Dev. |  | 0% | 5.391% | 4.279% | 3.873%
| PDM3 |  | -2.852% | 16.85% | -10.65% | -1.291%
| Z-Score (Absolute) |  | 0.1 | 2.22 | 1.25 | 0.1
| Analytical results for La in OREAS 464 (Certified Value 1.2 wt.%) | [8 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.175 | 1.185 | 1.194 | 1.211 | 1.219 | 1.22 | 1.194 | 1
| 1 | 2 | 1.195 | 1.185 | 1.185 | 1.185 | 1.219 | 1.22 | 1.177 | 0
| 1 | 3 | 1.195 | 1.194 | 1.219 | 1.194 | 1.219 | 1.22 | 1.177 | 16
| 1 | 4 | 1.195 | 1.177 | 1.194 | 1.211 | 1.219 | 1.23 | 1.177 | 1.1991
| 1 | 5 | 1.185 | 1.185 | 1.177 | 1.202 | 1.219 | 1.22 | 1.194 | 9
| 1 | 6 | 1.185 | 1.185 | 1.211 | 1.211 | 1.219 | 1.21 | 1.177
| Mean |  | 1.1883 | 1.1852 | 1.1967 | 1.2023 | 1.219 | 1.22 | 1.1827
| Median |  | 1.19 | 1.185 | 1.194 | 1.2065 | 1.219 | 1.22 | 1.177
| Std Dev. |  | 0.008165 | 0.0053821 | 0.015756 | 0.010912 | 2.4324e-16 | 0.0063246 | 0.0087788
| Rel.Std.Dev. |  | 0.6871% | 0.4541% | 1.317% | 0.9075% | 1.995e-14% | 0.5184% | 0.7423%
| PDM3 |  | -0.9015% | -1.166% | -0.2066% | 0.266% | 1.656% | 1.739% | -1.374%
| Z-Score (Absolute) |  | 0.49 | 0.67 | 0 | 0.34 | 1.35 | 1.39 | 0.84
| Analytical results for Lu in OREAS 464 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 18 | [5 empty] | 1
| 1 | 2 | < 90 | < 18 | [5 empty] | 1
| 1 | 3 | < 90 | < 18 | [5 empty] | 16
| 1 | 4 | < 90 | < 18 | [5 empty] | < 20
| 1 | 5 | < 90 | < 18 | [5 empty] | 7
| 1 | 6 | < 90 | < 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Mg in OREAS 464 (Indicative Value 0.855 wt.%) | [8 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.874 | 0.82 | 0.838 | 0.832 |  |  |  | 1
| 1 | 2 | 0.874 | 0.862 | 0.832 | 0.838 |  |  |  | 2
| 1 | 3 | 0.88 | 0.862 | 0.856 | 0.844 |  |  |  | 16
| 1 | 4 | 0.874 | 0.868 | 0.85 | 0.832 |  |  |  | 0.85509
| 1 | 5 | 0.868 | 0.85 | 0.85 | 0.856 |  |  |  | 8
| 1 | 6 | 0.874 | 0.844 | 0.874 | 0.862
| Mean |  | 0.874 | 0.851 | 0.85 | 0.844
| Median |  | 0.874 | 0.856 | 0.85 | 0.841
| Std Dev. |  | 0.0037947 | 0.017561 | 0.014697 | 0.012586
| Rel.Std.Dev. |  | 0.4342% | 2.064% | 1.729% | 1.491%
| PDM3 |  | 2.211% | -0.4788% | -0.5957% | -1.297%
| Z-Score (Absolute) |  | 4.53 | 0.1 | 0.1 | 1.25
| Analytical results for Mn in OREAS 464 (Indicative Value 0.306 wt.%) | [8 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.302 | 0.325 | 0.302 | 0.302 |  |  |  | 1
| 1 | 2 | 0.302 | 0.325 | 0.294 | 0.302 |  |  |  | 3
| 1 | 3 | 0.302 | 0.325 | 0.294 | 0.294 |  |  |  | 16
| 1 | 4 | 0.302 | 0.31 | 0.294 | 0.302 |  |  |  | 0.30592
| 1 | 5 | 0.302 | 0.325 | 0.294 | 0.302 |  |  |  | 9
| 1 | 6 | 0.302 | 0.325 | 0.294 | 0.302
| Mean |  | 0.302 | 0.3225 | 0.29533 | 0.30067
| Median |  | 0.302 | 0.325 | 0.294 | 0.302
| Std Dev. |  | 0 | 0.0061237 | 0.003266 | 0.003266
| Rel.Std.Dev. |  | 0% | 1.899% | 1.106% | 1.086%
| PDM3 |  | -1.28% | 5.421% | -3.46% | -1.716%
| Z-Score (Absolute) |  | 0.13 | 4.99 | 1.21 | 0.13
| Analytical results for Na in OREAS 464 (Indicative Value 0.174 wt.%) | [8 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.185 | 0.2 | 0.111 | 0.171 |  |  |  | 1
| 1 | 2 | 0.193 | 0.215 | 0.148 | 0.163 |  |  |  | 4
| 1 | 3 | 0.2 | 0.185 | 0.193 | 0.185 |  |  |  | 16
| 1 | 4 | 0.185 | 0.185 | 0.134 | 0.171 |  |  |  | 0.17396
| 1 | 5 | 0.178 | 0.2 | 0.171 | 0.171 |  |  |  | 10
| 1 | 6 | 0.185 | 0.185 | 0.104 | 0.171
| Mean |  | 0.18767 | 0.195 | 0.1435 | 0.172
| Median |  | 0.185 | 0.1925 | 0.141 | 0.171
| Std Dev. |  | 0.0076855 | 0.012247 | 0.034449 | 0.0071274
| Rel.Std.Dev. |  | 4.095% | 6.281% | 24.01% | 4.144%
| PDM3 |  | 7.878% | 12.09% | -17.51% | -1.128%
| Z-Score (Absolute) |  | 0.48 | 0.87 | 1.81 | 0.48
| Analytical results for Nb in OREAS 464 (Indicative Value 1938 ppm) | [8 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 2030 | 1890 | 1900 |  |  |  |  | 1
| 1 | 2 | 2030 | 1890 | 1900 |  |  |  |  | 5
| 1 | 3 | 2030 | 1890 | 1900 |  |  |  |  | 16
| 1 | 4 | 2030 | 1890 | 1900 |  |  |  |  | 1938.2384
| 1 | 5 | 2030 | 1890 | 1900 |  |  |  |  | 11
| 1 | 6 | 2030 | 1890 | 1900
| Mean |  | 2030 | 1890 | 1900
| Median |  | 2030 | 1890 | 1900
| Std Dev. |  | 0 | 0 | 2.4908e-13
| Rel.Std.Dev. |  | 0% | 0% | 1.311e-14%
| PDM3 |  | 4.734% | -2.489% | -1.973%
| Z-Score (Absolute) |  | 6.84 | 0.67 | 0
| Analytical results for Nd in OREAS 464 (Certified Value 0.964 wt.%) | [8 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 3
3
| 1 | 1 | 1.003 | 0.969 | 0.9 | 0.943 | 1.106 | 0.98 | 0.9602 | 1
| 1 | 2 | 1.012 | 0.969 | 0.9 | 0.943 | 1.115 | 0.99 | 0.9431 | 0
| 1 | 3 | 0.995 | 0.969 | 0.926 | 0.969 | 1.115 | 0.98 | 0.9431 | 16
| 1 | 4 | 1.012 | 0.969 | 0.917 | 0.977 | 1.115 | 0.99 | 0.9602 | 0.96367
| 1 | 5 | 1.003 | 0.969 | 0.934 | 0.952 | 1.097 | 0.99 | 0.9431 | 10
| 1 | 6 | 1.012 | 0.969 | 0.9 | 0.952 | 1.106 | 0.99 | 0.9602
| Mean |  | 1.0062 | 0.969 | 0.91283 | 0.956 | 1.109 | 0.98667 | 0.95165
| Median |  | 1.0075 | 0.969 | 0.9085 | 0.952 | 1.1105 | 0.99 | 0.95165
| Std Dev. |  | 0.0070261 | 1.2162e-16 | 0.015052 | 0.014 | 0.0073485 | 0.005164 | 0.0093661
| Rel.Std.Dev. |  | 0.6983% | 1.255e-14% | 1.649% | 1.464% | 0.6626% | 0.5234% | 0.9842%
| PDM3 |  | 4.409% | 0.5526% | -5.276% | -0.7964% | 15.08% | 2.386% | -1.248%
| Z-Score (Absolute) |  | 1.4 | 0 | 2.1 | 0.49 | 5.28 | 0.67 | 0.65
| Analytical results for Ni in OREAS 464 (Indicative Value 157 ppm) | [8 empty] | Indicative
Ni | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 160 | [6 empty] | 1
| 1 | 2 | 160 | [6 empty] | 6
| 1 | 3 | 160 | [6 empty] | 16
| 1 | 4 | 160 | [6 empty] | 157.1709
| 1 | 5 | 160 | [6 empty] | 12
| 1 | 6 | 160
| Mean |  | 160
| Median |  | 160
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 1.8%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 464 (Indicative Value 1.42 wt.%) | [8 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 09 | 14 | 15 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 2
2
| 1 | 1 | 1.4099 | 1.392 | 1.453 |  |  |  |  | 1
| 1 | 2 | 1.413 | 1.405 | 1.457 |  |  |  |  | 7
| 1 | 3 | 1.4117 | 1.401 | 1.444 |  |  |  |  | 16
| 1 | 4 | 1.4029 | 1.401 | 1.457 |  |  |  |  | 1.4202
| 1 | 5 | 1.4104 | 1.414 | 1.44 |  |  |  |  | 13
| 1 | 6 | 1.4095 | 1.388 | 1.453
| Mean |  | 1.4096 | 1.4002 | 1.4507
| Median |  | 1.4102 | 1.401 | 1.453
| Std Dev. |  | 0.0035098 | 0.0092826 | 0.0070616
| Rel.Std.Dev. |  | 0.249% | 0.663% | 0.4868%
| PDM3 |  | -0.7478% | -1.41% | 2.146%
| Z-Score (Absolute) |  | 0 | 0.67 | 2.93
| Analytical results for Pb in OREAS 464 (Indicative Value 278 ppm) | [8 empty] | Indicative
Pb | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 280 | [6 empty] | 1
| 1 | 2 | 280 | [6 empty] | 8
| 1 | 3 | 280 | [6 empty] | 16
| 1 | 4 | 280 | [6 empty] | 278.4998
| 1 | 5 | 280 | [6 empty] | 14
| 1 | 6 | 280
| Mean |  | 280
| Median |  | 280
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.5387%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 464 (Certified Value 2541 ppm) | [8 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 2480 | 2650 | 3080 | 2480 | 2910 | 2500 | 2400 | 1
| 1 | 2 | 2560 | 2570 | 3160 | 2480 | 2650 | 2500 | 2480 | 0
| 1 | 3 | 2480 | 2570 | 2820 | 2650 | 2650 | 2500 | 2400 | 16
| 1 | 4 | 2480 | 2650 | 3250 | 2570 | 2480 | 2600 | 2400 | 2540.6949
| 1 | 5 | 2560 | 2650 | 3080 | 2480 | 2650 | 2600 | 2480 | 11
| 1 | 6 | 2560 | 2570 | 2730 | 2400 | 2650 | 2600 | 2480
| Mean |  | 2520 | 2610 | 3020 | 2510 | 2665 | 2550 | 2440
| Median |  | 2520 | 2610 | 3080 | 2480 | 2650 | 2550 | 2440
| Std Dev. |  | 43.8178 | 43.8178 | 201.8911 | 87.178 | 137.9493 | 54.7723 | 43.8178
| Rel.Std.Dev. |  | 1.739% | 1.679% | 6.685% | 3.473% | 5.176% | 2.148% | 1.796%
| PDM3 |  | -0.8145% | 2.728% | 18.87% | -1.208% | 4.893% | 0.3662% | -3.963%
| Z-Score (Absolute) |  | 0.35 | 0.67 | 5.53 | 0.46 | 0.76 | 0 | 1.28
| Analytical results for Si in OREAS 464 (Indicative Value 9.35 wt.%) | [8 empty] | Indicative
Si | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 9.373 | 9.303 | 9.303 | 9.349 |  |  |  | 1
| 1 | 2 | 9.289 | 9.49 | 9.349 | 9.349 |  |  |  | 9
| 1 | 3 | 9.345 | 9.396 | 9.443 | 9.303 |  |  |  | 16
| 1 | 4 | 9.303 | 9.536 | 9.349 | 9.303 |  |  |  | 9.354
| 1 | 5 | 9.349 | 9.396 | 9.396 | 9.349 |  |  |  | 15
| 1 | 6 | 9.27 | 9.396 | 9.256 | 9.303
| Mean |  | 9.3215 | 9.4195 | 9.3493 | 9.326
| Median |  | 9.324 | 9.396 | 9.349 | 9.326
| Std Dev. |  | 0.040029 | 0.082185 | 0.066044 | 0.025195
| Rel.Std.Dev. |  | 0.4294% | 0.8725% | 0.7064% | 0.2702%
| PDM3 |  | -0.3471% | 0.7006% | -0.04951% | -0.299%
| Z-Score (Absolute) |  | 0.79 | 3.93 | 0.56 | 0.56
| Analytical results for Sm in OREAS 464 (Certified Value 1410 ppm) | [8 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 1380 | 1380 | > 4300 | 1380 | 1380 | 1600 | 1470 | 1
| 1 | 2 | 1380 | 1380 | > 4300 | 1470 | 1550 | 1600 | 1470 | 0
| 1 | 3 | 1380 | 1380 | > 4300 | 1380 | 1380 | 1600 | 1470 | 16
| 1 | 4 | 1380 | 1380 | > 4300 | 1380 | 1380 | 1600 | 1470 | 1409.7243
| 1 | 5 | 1380 | 1380 | > 4300 | 1380 | 1640 | 1600 | 1470 | 12
| 1 | 6 | 1380 | 1380 | > 4300 | 1470 | 1380 | 1600 | 1470
| Mean |  | 1380 | 1380 |  | 1410 | 1451.6667 | 1600 | 1470
| Median |  | 1380 | 1380 |  | 1380 | 1380 | 1600 | 1470
| Std Dev. |  | 0 | 0 |  | 46.4758 | 114.6153 | 2.4908e-13 | 0
| Rel.Std.Dev. |  | 0% | 0% |  | 3.296% | 7.895% | 1.557e-14% | 0%
| PDM3 |  | -2.109% | -2.109% |  | 0.01956% | 2.975% | 13.5% | 4.276%
| Z-Score (Absolute) |  | 0.67 | 0.68 | NA | 0.06 | 0.06 | 4.04 | 1.17
| Analytical results for Sn in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 160 | < 80 | [5 empty] | 1
| 1 | 2 | < 80 | < 80 | [5 empty] | 10
| 1 | 3 | < 80 | < 80 | [5 empty] | 16
| 1 | 4 | < 80 | < 80 | [5 empty] | < 100
| 1 | 5 | < 80 | < 80 | [5 empty] | 16
| 1 | 6 | 80 | < 80
| Mean |  | 120
| Median |  | 120
| Std Dev. |  | 56.5685
| Rel.Std.Dev. |  | 47.14%
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for S in OREAS 464 (Indicative Value 0.099 wt.%) | [8 empty] | Indicative
S | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.1005 | [6 empty] | 1
| 1 | 2 | 0.1013 | [6 empty] | 11
| 1 | 3 | 0.0969 | [6 empty] | 16
| 1 | 4 | 0.0941 | [6 empty] | 0.099378
| 1 | 5 | 0.1025 | [6 empty] | 17
| 1 | 6 | 0.1009
| Mean |  | 0.099367
| Median |  | 0.1007
| Std Dev. |  | 0.0031942
| Rel.Std.Dev. |  | 3.215%
| PDM3 |  | -0.01137%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 464 (Indicative Value 2001 ppm) | [8 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 2110 | 1860 | [5 empty] | 1
| 1 | 2 | 2200 | 1940 | [5 empty] | 1
| 1 | 3 | 2200 | 1690 | [5 empty] | 16
| 1 | 4 | 2110 | 1940 | [5 empty] | 2001.2402
| 1 | 5 | 2200 | 1690 | [5 empty] | 7
| 1 | 6 | 2200 | 1860
| Mean |  | 2170 | 1830
| Median |  | 2200 | 1860
| Std Dev. |  | 46.4758 | 114.1928
| Rel.Std.Dev. |  | 2.142% | 6.24%
| PDM3 |  | 8.433% | -8.557%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ta in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 410 | < 80 | [5 empty] | 1
| 1 | 2 | < 410 | < 80 | [5 empty] | 2
| 1 | 3 | < 410 | < 80 | [5 empty] | 16
| 1 | 4 | < 410 | < 80 | [5 empty] | < 100
| 1 | 5 | < 410 | < 80 | [5 empty] | 8
| 1 | 6 | < 410 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Tb in OREAS 464 (Indicative Value 49.4 ppm) | [8 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | 51 | [5 empty] | 1
| 1 | 2 | < 90 | 51 | [5 empty] | 3
| 1 | 3 | < 90 | 68 | [5 empty] | 16
| 1 | 4 | < 90 | 51 | [5 empty] | 49.4229
| 1 | 5 | < 90 | 51 | [5 empty] | 9
| 1 | 6 | 90 | 51
| Mean |  | 90 | 53.8333
| Median |  | 90 | 51
| Std Dev. |  |  | 6.9402
| Rel.Std.Dev. |  |  | 12.89%
| PDM3 |  | 82.1% | 8.924%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Th in OREAS 464 (Certified Value 506 ppm) | [8 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 480 | 530 | 440 | 530 | 598 | 490 |  | 1
| 1 | 2 | 490 | 530 | 530 | 440 | 642 | 500 |  | 0
| 1 | 3 | 480 | 620 | 350 | 530 | 615 | 490 |  | 16
| 1 | 4 | 490 | 530 | 350 | 530 | 606 | 500 |  | 506.0026
| 1 | 5 | 480 | 530 | 350 | 530 | 571 | 490 |  | 13
| 1 | 6 | 500 | 530 | 350 | 530 | 624 | 470
| Mean |  | 486.6667 | 545 | 395 | 515 | 609.3333 | 490
| Median |  | 485 | 530 | 350 | 530 | 610.5 | 490
| Std Dev. |  | 8.165 | 36.7423 | 75.2994 | 36.7423 | 24.1799 | 10.9545
| Rel.Std.Dev. |  | 1.678% | 6.742% | 19.06% | 7.134% | 3.968% | 2.236%
| PDM3 |  | -3.821% | 7.707% | -21.94% | 1.778% | 20.42% | -3.163%
| Z-Score (Absolute) |  | 0.36 | 0.99 | 2.58 | 0.28 | 2.64 | 0.28
| Analytical results for Ti in OREAS 464 (Indicative Value 1.97 wt.%) | [8 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.942 | 1.942 | 1.972 | 2.026 |  |  |  | 1
| 1 | 2 | 1.942 | 1.936 | 1.972 | 2.014 |  |  |  | 4
| 1 | 3 | 1.948 | 1.93 | 1.978 | 2.014 |  |  |  | 16
| 1 | 4 | 1.948 | 1.954 | 1.966 | 2.026 |  |  |  | 1.9706
| 1 | 5 | 1.954 | 1.942 | 2.002 | 2.014 |  |  |  | 10
| 1 | 6 | 1.936 | 1.942 | 1.972 | 2.014
| Mean |  | 1.945 | 1.941 | 1.977 | 2.018
| Median |  | 1.945 | 1.942 | 1.972 | 2.014
| Std Dev. |  | 0.0062929 | 0.007975 | 0.012822 | 0.0061968
| Rel.Std.Dev. |  | 0.3235% | 0.4109% | 0.6486% | 0.3071%
| PDM3 |  | -1.297% | -1.5% | 0.327% | 2.408%
| Z-Score (Absolute) |  | 0.6 | 0.75 | 0.6 | 2.14
| Analytical results for Tm in OREAS 464 (Indicative Value < 10 ppm) | [8 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 9 | [5 empty] | 1
| 1 | 2 | < 90 | < 9 | [5 empty] | 5
| 1 | 3 | < 90 | < 9 | [5 empty] | 16
| 1 | 4 | < 90 | < 9 | [5 empty] | < 10
| 1 | 5 | < 90 | < 9 | [5 empty] | 11
| 1 | 6 | < 90 | < 9
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 80 | < 250 | < 80 | 80 |  |  |  | 1
| 1 | 2 | < 80 | < 250 | < 80 | 80 |  |  |  | 6
| 1 | 3 | < 80 | < 250 | < 80 | < 80 |  |  |  | 16
| 1 | 4 | < 80 | < 250 | < 80 | < 80 |  |  |  | < 100
| 1 | 5 | < 80 | < 250 | < 80 | < 80 |  |  |  | 12
| 1 | 6 | < 80 | < 250 | < 80 | 80
| Mean |  |  |  |  | 80
| Median |  |  |  |  | 80
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for V in OREAS 464 (Indicative Value 258 ppm) | [8 empty] | Indicative
V | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 280 | 235 | [5 empty] | 1
| 1 | 2 | 280 | 230 | [5 empty] | 7
| 1 | 3 | 280 | 218 | [5 empty] | 16
| 1 | 4 | 280 | 252 | [5 empty] | 257.6742
| 1 | 5 | 280 | 241 | [5 empty] | 13
| 1 | 6 | 280 | 235
| Mean |  | 280 | 235.1667
| Median |  | 280 | 235
| Std Dev. |  | 0 | 11.3034
| Rel.Std.Dev. |  | 0% | 4.807%
| PDM3 |  | 8.664% | -8.735%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for W in OREAS 464 (Indicative Value < 100 ppm) | [8 empty] | Indicative
W | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 8
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 14
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 464 (Certified Value 442 ppm) | [8 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 300 | 550 | 470 | 470 | 550 | 400 | 402 | 1
| 1 | 2 | 300 | 390 | 390 | 390 | 470 | 400 | 402 | 0
| 1 | 3 | 300 | 470 | 390 | 390 | 550 | 500 | 370 | 16
| 1 | 4 | 300 | 470 | 470 | 390 | 470 | 400 | 386 | 442.4618
| 1 | 5 | 300 | 470 | 470 | 470 | 550 | 400 | 386 | 14
| 1 | 6 | 300 | 470 | 470 | 470 | 470 | 400 | 402
| Mean |  | 300 | 470 | 443.3333 | 430 | 510 | 416.6667 | 391.3333
| Median |  | 300 | 470 | 470 | 430 | 510 | 400 | 394
| Std Dev. |  | 6.2269e-14 | 50.5964 | 41.3118 | 43.8178 | 43.8178 | 40.8248 | 13.0639
| Rel.Std.Dev. |  | 2.076e-14% | 10.77% | 9.318% | 10.19% | 8.592% | 9.798% | 3.338%
| PDM3 |  | -32.2% | 6.224% | 0.197% | -2.816% | 15.26% | -5.83% | -11.56%
| Z-Score (Absolute) |  | 2.28 | 0.67 | 0.22 | 0 | 1.35 | 0.57 | 0.72
| Analytical results for Yb in OREAS 464 (Indicative Value 17.6 ppm) | [8 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 1
1
| 1 | 1 | < 90 | < 90 | 18 |  |  |  |  | 1
| 1 | 2 | < 90 | 90 | 18 |  |  |  |  | 9
| 1 | 3 | < 90 | 90 | 18 |  |  |  |  | 16
| 1 | 4 | < 90 | 90 | 18 |  |  |  |  | 17.5639
| 1 | 5 | < 90 | 90 | 18 |  |  |  |  | 15
| 1 | 6 | < 90 | < 90 | 18
| Mean |  |  | 90 | 18
| Median |  |  | 90 | 18
| Std Dev. |  |  | 0 | 0
| Rel.Std.Dev. |  |  | 0% | 0%
| PDM3 |  |  | 412.4% | 2.483%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Zn in OREAS 464 (Indicative Value 1259 ppm) | [8 empty] | Indicative
Zn | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 1210 | [6 empty] | 1
| 1 | 2 | 1290 | [6 empty] | 10
| 1 | 3 | 1290 | [6 empty] | 16
| 1 | 4 | 1290 | [6 empty] | 1258.569
| 1 | 5 | 1290 | [6 empty] | 16
| 1 | 6 | 1210
| Mean |  | 1263.3333
| Median |  | 1290
| Std Dev. |  | 41.3118
| Rel.Std.Dev. |  | 3.27%
| PDM3 |  | 0.3786%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 464 (Indicative Value 457 ppm) | [8 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 440 | 440 | [5 empty] | 1
| 1 | 2 | 520 | 440 | [5 empty] | 11
| 1 | 3 | 520 | 440 | [5 empty] | 16
| 1 | 4 | 520 | 370 | [5 empty] | 456.5196
| 1 | 5 | 440 | 440 | [5 empty] | 17
| 1 | 6 | 440 | 440
| Mean |  | 480 | 428.3333
| Median |  | 480 | 440
| Std Dev. |  | 43.8178 | 28.5774
| Rel.Std.Dev. |  | 9.129% | 6.672%
| PDM3 |  | 5.143% | -6.174%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 464 (Indicative Value 0.503 wt.%) | [8 empty] | Indicative
H2O- | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | at 1000°C | [6 empty] | 3
3
| 1 | 1 | 0.51 | [6 empty] | 1
| 1 | 2 | 0.49 | [6 empty] | 13
| 1 | 3 | 0.52 | [6 empty] | 16
| 1 | 4 | 0.5 | [6 empty] | 0.50333
| 1 | 5 | 0.5 | [6 empty] | 19
| 1 | 6 | 0.5
| Mean |  | 0.50333
| Median |  | 0.5
| Std Dev. |  | 0.010328
| Rel.Std.Dev. |  | 2.052%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI in OREAS 464 (Certified Value 0.697 wt.%) | [8 empty] | Certified
LOI | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 18 | wt.%
|  |  | at 1000°C | at 1000°C | at 1000°C | BF*XRF | at 1000°C | BF*XRF | at 1000°C | 3
3
| 1 | 1 | 0.78 | 0.64 | 0.62 | 0.8 | 1.06 | 0.632 | 1.1791 | 1
| 1 | 2 | 0.71 | 0.6 | 0.68 | 0.71 | 1.07 | 0.665 | 1.1395 | 0
| 1 | 3 | 0.81 | 0.63 | 0.7 | 0.81 | 1.04 | 0.639 | 1.1588 | 16
| 1 | 4 | 0.75 | 0.58 | 0.67 | 0.75 | 1.08 | 0.653 | 1.0585 | 0.6969
| 1 | 5 | 0.79 | 0.6 | 0.66 | 0.8 | 1.06 | 0.651 | 1.0598 | 16
| 1 | 6 | 0.79 | 0.62 | 0.68 | 0.75 | 1.08 | 0.679 | 1.1391
| Mean |  | 0.77167 | 0.61167 | 0.66833 | 0.77 | 1.065 | 0.65317 | 1.1225
| Median |  | 0.785 | 0.61 | 0.675 | 0.775 | 1.065 | 0.652 | 1.1393
| Std Dev. |  | 0.036009 | 0.022286 | 0.027142 | 0.039497 | 0.015166 | 0.017093 | 0.051208
| Rel.Std.Dev. |  | 4.666% | 3.643% | 4.061% | 5.129% | 1.424% | 2.617% | 4.562%
| PDM3 |  | 10.73% | -12.23% | -4.099% | 10.49% | 52.82% | -6.275% | 61.07%
| Z-Score (Absolute) |  | 0.01 | 0.91 | 0.53 | 0 | 1.7 | 0.67 | 2.03
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 464 (Indicative Value 10.4 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 20 | 22 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | [12 empty] | 1
1
| 1 | 1 | < 5 | 12 | 5 | [12 empty] | 1
| 1 | 2 | < 5 | 18 | 5 | [12 empty] | 15
| 1 | 3 | < 5 | 18 | 5 | [12 empty] | 16
| 1 | 4 | < 5 | 12 | 5 | [12 empty] | 10.4167
| 1 | 5 | < 5 | 17 | 5 | [12 empty] | 21
| 1 | 6 | < 5 | 18 | 5
| Mean |  |  | 15.8333 | 5
| Median |  |  | 17.5 | 5
| Std Dev. |  |  | 2.9944 | 0
| Rel.Std.Dev. |  |  | 18.91% | 0%
| PDM3 |  |  | 52% | -52%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Al in OREAS 464 (Certified Value 4.63 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 4.752 | 4.56 | 4.49 | 4.5323 | N.A. | 4.4318 | 4.329 | 4.99 | 4.71 | 4.76 | [5 empty] | 1
| 1 | 2 | 4.712 | 4.62 | 4.56 | 4.4691 | N.A. | 4.4867 | 4.488 | 4.83 | 4.68 | 4.74 | [5 empty] | 1
| 1 | 3 | 4.85 | 4.6 | 4.55 | 4.5409 | N.A. | 4.4731 | 4.435 | 4.92 | 4.82 | 4.78 | [5 empty] | 16
| 1 | 4 | 4.715 | 4.59 | 4.55 | 4.4416 | N.A. | 4.5314 | 4.456 | 4.82 | 4.75 | 4.75 | [5 empty] | 4.6305
| 1 | 5 | 4.755 | 4.56 | 4.61 | 4.4135 | N.A. | 4.5212 | 4.578 | 4.73 | 4.85 | 4.79 | [5 empty] | 18
| 1 | 6 | 4.737 | 4.47 | 4.56 | 4.4456 | N.A. | 4.5246 | 4.536 | 4.74 | 4.79 | 4.72
| Mean |  | 4.7535 | 4.5667 | 4.5533 | 4.4739 |  | 4.4948 | 4.4703 | 4.8383 | 4.7667 | 4.7567
| Median |  | 4.7445 | 4.575 | 4.555 | 4.4573 |  | 4.504 | 4.472 | 4.825 | 4.77 | 4.755
| Std Dev. |  | 0.050584 | 0.052789 | 0.038297 | 0.0518 |  | 0.038577 | 0.08677 | 0.10147 | 0.06532 | 0.02582
| Rel.Std.Dev. |  | 1.064% | 1.156% | 0.8411% | 1.158% |  | 0.8583% | 1.941% | 2.097% | 1.37% | 0.5428%
| PDM3 |  | 2.657% | -1.378% | -1.666% | -3.382% |  | -2.93% | -3.458% | 4.489% | 2.942% | 2.726%
| Z-Score (Absolute) |  | 1.31 | 0 | 0.09 | 0.65 | NA | 0.5 | 0.67 | 1.9 | 1.4 | 1.33
| Analytical results for As in OREAS 464 (Indicative Value 252 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 42 | < 100 | 320.9894 | 330 | 20 | 570 | [9 empty] | 1
| 1 | 2 | 35 | < 100 | 331.8848 | 310 | 20 | 550 | [9 empty] | 16
| 1 | 3 | 38 | < 100 | 323.7574 | 310 | < 10 | 540 | [9 empty] | 16
| 1 | 4 | 40 | < 100 | 321.9192 | 320 | 40 | 570 | [9 empty] | 251.8708
| 1 | 5 | 41 | < 100 | 326.7707 | 310 | 40 | 550 | [9 empty] | 22
| 1 | 6 | 36 | < 100 | 323.8037 | 290 | 50 | 550
| Mean |  | 38.6667 |  | 324.8542 | 311.6667 | 34 | 555
| Median |  | 39 |  | 323.7806 | 310 | 40 | 550
| Std Dev. |  | 2.8048 |  | 3.9739 | 13.2916 | 13.4164 | 12.2474
| Rel.Std.Dev. |  | 7.254% |  | 1.223% | 4.265% | 39.46% | 2.207%
| PDM3 |  | -84.65% |  | 28.98% | 23.74% | -86.5% | 120.4%
| Z-Score (Absolute) |  | 0.76 | NA | 0.04 | 0 | 0.78 | 0.67
| Analytical results for B in OREAS 464 (Indicative Value 108 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*OES | PF*OES | [13 empty] | 0
0
| 1 | 1 | 112 | 100 | [13 empty] | 1
| 1 | 2 | 113 | 100 | [13 empty] | 17
| 1 | 3 | 112 | 120 | [13 empty] | 16
| 1 | 4 | 121 | 80 | [13 empty] | 107.5833
| 1 | 5 | 117 | 100 | [13 empty] | 23
| 1 | 6 | 116 | 100
| Mean |  | 115.1667 | 100
| Median |  | 114.5 | 100
| Std Dev. |  | 3.5449 | 12.6491
| Rel.Std.Dev. |  | 3.078% | 12.65%
| PDM3 |  | 7.049% | -7.049%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ba in OREAS 464 (Certified Value 1657 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1705 | 1555 | 1715 | 1725 | 1660 | 1534.1731 | 1588 | 1580.2 | 1383.224 | N.A. | 1700 | 1623.3857 | 1630 | 1720 | 1360 | 1
| 1 | 2 | 1645 | 1575 | 1730 | 1705 | 1730 | 1536.1307 | 1630 | 1610.4 | 1367 | N.A. | 1700 | 1658.4784 | 1650 | 1730 | 1330 | 2
| 1 | 3 | 1705 | 1590 | 1695 | 1750 | 1660 | 1576.5297 | 1601 | 1619.6 | 1381.904 | N.A. | 1700 | 1613.1219 | 1700 | 1700 | 1490 | 16
| 1 | 4 | 1735 | 1710 | 1745 | 1735 | 1690 | 1505.1566 | 1628 | 1601.6 | 1370.792 | N.A. | 1700 | 1629.4235 | 1570 | 1690 | 1360 | 1656.7244
| 1 | 5 | 1695 | 1580 | 1780 | 1765 | 1775 | 1540.5455 | 1626 | 1639.4 | 1366.592 | N.A. | 1700 | 1667.775 | 1650 | 1680 | 1370 | 19
| 1 | 6 | 1675 | 1500 | 1765 | 1745 | 1705 | 1525.0957 | 1619 | 1633.2 | 1382.432 | N.A. | 1700 | 1622.9397 | 1620 | 1740 | 1370
| Mean |  | 1693.3333 | 1585 | 1738.3333 | 1737.5 | 1703.3333 | 1536.2719 | 1615.3333 | 1614.0667 | 1375.324 |  | 1700 | 1635.854 | 1636.6667 | 1710 | 1380
| Median |  | 1700 | 1577.5 | 1737.5 | 1740 | 1697.5 | 1535.1519 | 1622.5 | 1615 | 1376.348 |  | 1700 | 1626.4046 | 1640 | 1710 | 1365
| Std Dev. |  | 30.605 | 69.1375 | 31.57 | 20.9165 | 44.2342 | 23.3942 | 17.0607 | 21.6993 | 8.0288 |  | 0 | 21.9597 | 42.7395 | 23.6643 | 55.857
| Rel.Std.Dev. |  | 1.807% | 4.362% | 1.816% | 1.204% | 2.597% | 1.523% | 1.056% | 1.344% | 0.5838% |  | 0% | 1.342% | 2.611% | 1.384% | 4.048%
| PDM3 |  | 2.21% | -4.329% | 4.926% | 4.876% | 2.813% | -7.271% | -2.498% | -2.575% | -16.99% |  | 2.612% | -1.26% | -1.211% | 3.216% | -16.7%
| Z-Score (Absolute) |  | 0.55 | 0.73 | 0.98 | 0.97 | 0.64 | 0.96 | 0.2 | 0.21 | 2.5 | NA | 0.61 | 0 | 0 | 0.71 | 2.67
| Analytical results for Be in OREAS 464 (Certified Value 15.6 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 17 | 22 | 23 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 1
1
| 1 | 1 | 16.1984 | 16 | 14.4 | 10.3927 | 17 | 16 | 11 | [8 empty] | 1
| 1 | 2 | 15.5972 | 18 | 14.4 | 10.102 | 17 | 17 | 13 | [8 empty] | 0
| 1 | 3 | 16.009 | 17 | 14.6 | 10.1381 | 15 | 18 | 14 | [8 empty] | 16
| 1 | 4 | 16.8477 | 16 | 14.6 | 10.1341 | 16 | 15 | 13 | [8 empty] | 15.6289
| 1 | 5 | 16.701 | 18 | 15.1 | 10.2156 | 16 | 16 | 13 | [8 empty] | 20
| 1 | 6 | 16.566 | 17 | 15.7 | 10.1306 | 17 | 16 | 15
| Mean |  | 16.3199 | 17 | 14.8 | 10.1855 | 16.3333 | 16.3333 | 13.1667
| Median |  | 16.3822 | 17 | 14.6 | 10.1361 | 16.5 | 16 | 13
| Std Dev. |  | 0.47287 | 0.89443 | 0.5099 | 0.10836 | 0.8165 | 1.0328 | 1.3292
| Rel.Std.Dev. |  | 2.897% | 5.261% | 3.445% | 1.064% | 4.999% | 6.323% | 10.09%
| PDM3 |  | 4.421% | 8.773% | -5.303% | -34.83% | 4.507% | 4.507% | -15.75%
| Z-Score (Absolute) |  | 0 | 0.67 | 1.69 | 6.08 | 0.01 | 0.01 | 3.13
| Analytical results for Bi in OREAS 464 (Certified Value 3.8 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 4.21 | 4.2 | 3.5 | 1.9 | 4 | 3.8 | [9 empty] | 1
| 1 | 2 | 4.3465 | 4.2 | 3.5 | 2.1 | 3 | 3.4 | [9 empty] | 0
| 1 | 3 | 4.0855 | 5.1 | 4 | 2.1 | 3 | 3.4 | [9 empty] | 16
| 1 | 4 | 3.9479 | 4.2 | 4 | 2 | 3 | 3.6 | [9 empty] | 3.7965
| 1 | 5 | 4.9774 | 4.2 | 3.5 | 2.1 | 3 | 3.5 | [9 empty] | 21
| 1 | 6 | 4.0722 | 4.2 | 4 | 2.1 | 4 | 3.7
| Mean |  | 4.2732 | 4.35 | 3.75 | 2.05 | 3.3333 | 3.5667
| Median |  | 4.1477 | 4.2 | 3.75 | 2.1 | 3 | 3.55
| Std Dev. |  | 0.37062 | 0.36742 | 0.27386 | 0.083666 | 0.5164 | 0.1633
| Rel.Std.Dev. |  | 8.673% | 8.447% | 7.303% | 4.081% | 15.49% | 4.578%
| PDM3 |  | 12.56% | 14.58% | -1.224% | -46% | -12.2% | -6.053%
| Z-Score (Absolute) |  | 0.8 | 0.91 | 0.15 | 2.71 | 0.55 | 0.15
| Analytical results for Ca in OREAS 464 (Certified Value 0.896 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.854 | 0.9 | 0.98 | 0.79801 | N.A. | 0.89 | 0.86419 | 0.893 | 0.967 | 0.9 | 0.9 |  |  |  |  | 1
| 1 | 2 | 0.828 | 0.9 | 0.94 | 0.81277 | N.A. | 0.91 | 0.8599 | 0.893 | 0.929 | 0.9 | 0.9 |  |  |  |  | 3
| 1 | 3 | 0.872 | 0.9 | 0.92 | 0.78896 | N.A. | 1.01 | 0.86276 | 0.901 | 0.957 | 0.9 | 0.9 |  |  |  |  | 16
| 1 | 4 | 0.824 | 0.9 | 0.94 | 0.8192 | N.A. | 0.86 | 0.85633 | 0.901 | 0.923 | 0.9 | 0.9 |  |  |  |  | 0.89604
| 1 | 5 | 0.883 | 0.9 | 0.93 | 0.79253 | N.A. | 0.86 | 0.86705 | 0.893 | 0.909 | 0.9 | 0.9 |  |  |  |  | 22
| 1 | 6 | 0.912 | 0.9 | 0.94 | 0.78975 | N.A. | 0.9 | 0.86705 | 0.908 | 0.908 | 0.9 | 0.8
| Mean |  | 0.86217 | 0.9 | 0.94167 | 0.8002 |  | 0.905 | 0.86288 | 0.89817 | 0.93217 | 0.9 | 0.88333
| Median |  | 0.863 | 0.9 | 0.94 | 0.79527 |  | 0.895 | 0.86347 | 0.897 | 0.926 | 0.9 | 0.9
| Std Dev. |  | 0.033778 | 1.2162e-16 | 0.020412 | 0.012792 |  | 0.055408 | 0.0042025 | 0.0062102 | 0.024677 | 1.2162e-16 | 0.040825
| Rel.Std.Dev. |  | 3.918% | 1.351e-14% | 2.168% | 1.599% |  | 6.122% | 0.487% | 0.6914% | 2.647% | 1.351e-14% | 4.622%
| PDM3 |  | -3.78% | 0.4421% | 5.092% | -10.7% |  | 1% | -3.701% | 0.2375% | 4.032% | 0.4421% | -1.418%
| Z-Score (Absolute) |  | 1.05 | 0.32 | 1.84 | 3.3 | NA | 0.26 | 1.02 | 0.26 | 1.49 | 0.32 | 0.28
| Analytical results for Cd in OREAS 464 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | < 1 | < 1 | < 10 | 0.5 | [11 empty] | 1
| 1 | 2 | < 1 | < 1 | < 10 | 0.6 | [11 empty] | 18
| 1 | 3 | < 1 | < 1 | < 10 | 0.4 | [11 empty] | 16
| 1 | 4 | < 1 | < 1 | < 10 | 0.6 | [11 empty] | < 1
| 1 | 5 | < 1 | < 1 | < 10 | 0.8 | [11 empty] | 24
| 1 | 6 | < 1 | < 1 | < 10 | 0.7
| Mean |  |  |  |  | 0.6
| Median |  |  |  |  | 0.6
| Std Dev. |  |  |  |  | 0.14142
| Rel.Std.Dev. |  |  |  |  | 23.57%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for Ce in OREAS 464 (Certified Value 1.53 wt.%) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.59 | 1.515 | 1.535 | > 1 | 1.47 | 1.4487 | 1.523 | 1.4915 | 1.3984 | 1.547 | 1.5968 | 1.6898 | 1.58 | 1.54 | 1.7 | 1
| 1 | 2 | 1.565 | 1.565 | 1.52 | > 1 | 1.485 | 1.4387 | 1.568 | 1.5019 | 1.4503 | 1.555 | 1.6096 | 1.6894 | 1.55 | 1.52 | 1.67 | 0
| 1 | 3 | 1.625 | 1.61 | 1.455 | > 1 | 1.5 | 1.4857 | 1.5258 | 1.5025 | 1.4365 | 1.555 | 1.5865 | 1.6578 | 1.59 | 1.53 | 1.65 | 16
| 1 | 4 | 1.63 | 1.545 | 1.555 | > 1 | 1.5 | 1.4282 | 1.5303 | 1.5076 | 1.4225 | 1.506 | 1.5882 | 1.6602 | 1.54 | 1.56 | 1.69 | 1.5274
| 1 | 5 | 1.57 | 1.545 | 1.5 | > 1 | 1.48 | 1.4583 | 1.5418 | 1.5239 | 1.4784 | 1.53 | 1.6192 | 1.6991 | 1.59 | 1.51 | 1.74 | 23
| 1 | 6 | 1.58 | 1.555 | 1.515 | > 1 | 1.56 | 1.4596 | 1.5408 | 1.5092 | 1.4795 | 1.522 | 1.6278 | 1.6967 | 1.53 | 1.54 | 1.72
| Mean |  | 1.5933 | 1.5558 | 1.5133 |  | 1.4992 | 1.4532 | 1.5383 | 1.5061 | 1.4443 | 1.5358 | 1.6047 | 1.6822 | 1.5633 | 1.5333 | 1.695
| Median |  | 1.585 | 1.55 | 1.5175 |  | 1.4925 | 1.4535 | 1.5355 | 1.505 | 1.4434 | 1.5385 | 1.6032 | 1.6896 | 1.565 | 1.535 | 1.695
| Std Dev. |  | 0.027869 | 0.031371 | 0.034157 |  | 0.032003 | 0.019868 | 0.016495 | 0.0107 | 0.031895 | 0.019874 | 0.016934 | 0.018376 | 0.026583 | 0.017512 | 0.032711
| Rel.Std.Dev. |  | 1.749% | 2.016% | 2.257% |  | 2.135% | 1.367% | 1.072% | 0.7105% | 2.208% | 1.294% | 1.055% | 1.092% | 1.7% | 1.142% | 1.93%
| PDM3 |  | 4.318% | 1.863% | -0.9198% |  | -1.847% | -4.856% | 0.7136% | -1.393% | -5.441% | 0.5533% | 5.061% | 10.13% | 2.354% | 0.3896% | 10.97%
| Z-Score (Absolute) |  | 0.94 | 0.31 | 0.4 | NA | 0.83 | 1.4 | 0.02 | 0.52 | 1.54 | 0.02 | 1.12 | 2.41 | 0.44 | 0.06 | 2.63
| Analytical results for Co in OREAS 464 (Certified Value 18 ppm) | [16 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [6 empty] | 1
1
| 1 | 1 | 19.8 | 20.8463 | N.A. | < 50 | 23.4654 | 5.4 | < 10 | 40 | 16.2 | [6 empty] | 1
| 1 | 2 | 19.4 | 21.1559 | N.A. | < 50 | 16.2513 | 18.7 | < 10 | 20 | 14.8 | [6 empty] | 0
| 1 | 3 | 19.9 | 21.1965 | N.A. | < 50 | 16.9955 | 18.7 | < 10 | 20 | 14.5 | [6 empty] | 16
| 1 | 4 | 19.4 | 20.5674 | N.A. | < 50 | 16.5757 | 4.4 | < 10 | 20 | 15.3 | [6 empty] | 18.0159
| 1 | 5 | 20.7 | 22.5026 | N.A. | < 50 | 19.2961 | 18.5 | < 10 | 40 | 15.6 | [6 empty] | 24
| 1 | 6 | 20.1 | 22.8195 | N.A. | < 50 | 17.8555 | 17.8 | < 10 | 20 | 16.4
| Mean |  | 19.8833 | 21.5147 |  |  | 18.4066 | 13.9167 |  | 26.6667 | 15.4667
| Median |  | 19.85 | 21.1762 |  |  | 17.4255 | 18.15 |  | 20 | 15.45
| Std Dev. |  | 0.48751 | 0.92223 |  |  | 2.7083 | 6.9993 |  | 10.328 | 0.75277
| Rel.Std.Dev. |  | 2.452% | 4.287% |  |  | 14.71% | 50.29% |  | 38.73% | 4.867%
| PDM3 |  | 10.37% | 19.42% |  |  | 2.169% | -22.75% |  | 48.02% | -14.15%
| Z-Score (Absolute) |  | 0.29 | 0.68 | NA | NA | 0.29 | 0.67 | NA | 1.9 | 0.75
| Analytical results for Cr in OREAS 464 (Certified Value 395 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES | PF*OES/MS | 0
0
| 1 | 1 | 410 | 370 | 410 | 420 | 350 | < 1000 | 375 | 377.423 | N.A. | 400 | 405.3149 | 200 | 720 | 290 | 410 | 1
| 1 | 2 | 390 | 360 | 410 | 420 | 370 | < 1000 | 371 | 377.595 | N.A. | 400 | 411.0432 | 300 | 710 | 280 | 410 | 4
| 1 | 3 | 410 | 370 | 410 | 430 | 350 | < 1000 | 358 | 375.389 | N.A. | 400 | 403.1851 | 300 | 800 | 300 | 400 | 16
| 1 | 4 | 410 | 410 | 410 | 420 | 360 | < 1000 | 400 | 376.386 | N.A. | 400 | 408.7472 | 200 | 700 | 300 | 400 | 395.1231
| 1 | 5 | 410 | 370 | 440 | 420 | 410 | < 1000 | 387 | 377.12 | N.A. | 400 | 407.6887 | 400 | 570 | 300 | 410 | 25
| 1 | 6 | 400 | 330 | 420 | 410 | 400 | < 1000 | 366 | 380.699 | N.A. | 400 | 419.7944 | 400 | 760 | 300 | 410
| Mean |  | 405 | 368.3333 | 416.6667 | 420 | 373.3333 |  | 376.1667 | 377.4353 |  | 400 | 409.2956 | 300 | 710 | 295 | 406.6667
| Median |  | 410 | 370 | 410 | 420 | 365 |  | 373 | 377.2715 |  | 400 | 408.218 | 300 | 715 | 300 | 410
| Std Dev. |  | 8.3666 | 25.6255 | 12.1106 | 6.3246 | 25.8199 |  | 15.1449 | 1.7925 |  | 6.2269e-14 | 5.8197 | 89.4427 | 77.9744 | 8.3666 | 5.164
| Rel.Std.Dev. |  | 2.066% | 6.957% | 2.907% | 1.506% | 6.916% |  | 4.026% | 0.4749% |  | 1.557e-14% | 1.422% | 29.81% | 10.98% | 2.836% | 1.27%
| PDM3 |  | 2.5% | -6.78% | 5.452% | 6.296% | -5.515% |  | -4.798% | -4.477% |  | 1.234% | 3.587% | -24.07% | 79.69% | -25.34% | 2.922%
| Z-Score (Absolute) |  | 0.24 | 0.95 | 0.36 | 0.6 | 0.8 | NA | 0.71 | 0.67 | NA | 0 | 0.28 | 2.99 | 9.26 | 3.05 | 0.2
| Analytical results for Cs in OREAS 464 (Certified Value 0.27 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 2
2
| 1 | 1 | 0.29 | 0.22 | 0.31 | 0.24 | 0.29 | 0.26114 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 1
| 1 | 2 | 0.33 | 0.18 | 0.29 | 0.26 | 0.29 | 0.3387 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 0
| 1 | 3 | 0.32 | 0.22 | 0.25 | 0.27 | 0.29 | 0.2876 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 16
| 1 | 4 | 0.31 | 0.28 | 0.31 | 0.24 | 0.27 | 0.23432 | 0.3 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5 |  |  | 0.27145
| 1 | 5 | 0.37 | 0.21 | 0.3 | 0.26 | 0.38 | 0.27358 | 0.2 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.6 |  |  | 26
| 1 | 6 | 0.31 | 0.22 | 0.29 | 0.22 | 0.32 | 0.22366 | 0.2 | 0.4 | N.A. | < 1 | < 0.1 | < 1 | 0.5
| Mean |  | 0.32167 | 0.22167 | 0.29167 | 0.24833 | 0.30667 | 0.26983 | 0.26667 | 0.4 |  |  |  |  | 0.51667
| Median |  | 0.315 | 0.22 | 0.295 | 0.25 | 0.29 | 0.26736 | 0.3 | 0.4 |  |  |  |  | 0.5
| Std Dev. |  | 0.027142 | 0.032506 | 0.022286 | 0.018348 | 0.039328 | 0.04133 | 0.05164 | 6.0809e-17 |  |  |  |  | 0.040825
| Rel.Std.Dev. |  | 8.438% | 14.66% | 7.641% | 7.389% | 12.82% | 15.32% | 19.36% | 1.52e-14% |  |  |  |  | 7.902%
| PDM3 |  | 18.5% | -18.34% | 7.447% | -8.517% | 12.97% | -0.5965% | -1.763% | 47.36% |  |  |  |  | 90.33%
| Z-Score (Absolute) |  | 0.67 | 1.84 | 0 | 0.97 | 0.01 | 0.49 | 0.56 | 2.44 | NA | NA | NA | NA | 4.68
| Analytical results for Cu in OREAS 464 (Indicative Value 104 ppm) | [16 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | N.A. | < 100 | 109.2261 | 75 | < 80 | 100 | 110 | [8 empty] | 1
| 1 | 2 | N.A. | < 100 | 113.029 | 364 | < 80 | 100 | 110 | [8 empty] | 19
| 1 | 3 | N.A. | < 100 | 105.7489 | 100 | < 80 | 100 | 110 | [8 empty] | 16
| 1 | 4 | N.A. | < 100 | 107.6508 | 82 | < 80 | 100 | 120 | [8 empty] | 103.6968
| 1 | 5 | N.A. | < 100 | 106.6592 | 102 | < 80 | 100 | 110 | [8 empty] | 25
| 1 | 6 | N.A. | < 100 | 108.8084 | 114 | < 80 | 100 | 110
| Mean |  |  |  | 108.5204 | 139.5 |  | 100 | 111.6667
| Median |  |  |  | 108.2296 | 101 |  | 100 | 110
| Std Dev. |  |  |  | 2.5616 | 110.8905 |  | 0 | 4.0825
| Rel.Std.Dev. |  |  |  | 2.361% | 79.49% |  | 0% | 3.656%
| PDM3 |  |  |  | 4.652% | 34.53% |  | -3.565% | 7.686%
| Z-Score (Absolute) |  | NA | NA | 0.49 | 1.12 | NA | 0.49 | 0.86
| Analytical results for Dy in OREAS 464 (Certified Value 178 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 173 | 177 | 180 | 180 | 172 | 187.7093 | 181.5 | 178.4 | 140.8564 | 170 | 185.5 | 160.2901 | 185 | 188 | 166 | 1
| 1 | 2 | 174 | 174 | 181 | 181 | 174 | 186.5423 | 181.9 | 176.4 | 141.6618 | 170 | 185 | 162.2125 | 231 | 190 | 168 | 5
| 1 | 3 | 174.5 | 178 | 179.5 | 182.5 | 182.5 | 195.1772 | 179.8 | 177.1 | 138.9395 | 170 | 182 | 158.822 | 231 | 182 | 167 | 16
| 1 | 4 | 176 | 178.5 | 192 | 183.5 | 174.5 | 186.3525 | 180.5 | 180.2 | 140.3145 | 170 | 179 | 157.5899 | 181 | 186 | 170 | 178.088
| 1 | 5 | 168.5 | 175.5 | 186.5 | 184 | 170.5 | 189.8399 | 180 | 184.8 | 143.772 | 170 | 185.5 | 160.9251 | 236 | 182 | 175 | 27
| 1 | 6 | 174 | 175.5 | 184 | 179.5 | 179 | 191.668 | 183.3 | 178.8 | 141.5039 | 170 | 185 | 159.1883 | 229 | 187 | 174
| Mean |  | 173.3333 | 176.4167 | 183.8333 | 181.75 | 175.4167 | 189.5482 | 181.1667 | 179.2833 | 141.1747 | 170 | 183.6667 | 159.838 | 215.5 | 185.8333 | 170
| Median |  | 174 | 176.25 | 182.5 | 181.75 | 174.25 | 188.7746 | 181 | 178.6 | 141.1802 | 170 | 185 | 159.7392 | 230 | 186.5 | 169
| Std Dev. |  | 2.5626 | 1.7151 | 4.8028 | 1.8641 | 4.5102 | 3.432 | 1.3322 | 3.0123 | 1.6081 | 0 | 2.6394 | 1.6451 | 25.3121 | 3.2506 | 3.7417
| Rel.Std.Dev. |  | 1.478% | 0.9722% | 2.613% | 1.026% | 2.571% | 1.811% | 0.7353% | 1.68% | 1.139% | 0% | 1.437% | 1.029% | 11.75% | 1.749% | 2.201%
| PDM3 |  | -2.67% | -0.9385% | 3.226% | 2.056% | -1.5% | 6.435% | 1.729% | 0.6712% | -20.73% | -4.542% | 3.133% | -10.25% | 21.01% | 4.349% | -4.542%
| Z-Score (Absolute) |  | 0.46 | 0.21 | 0.67 | 0.43 | 0.33 | 1.36 | 0.36 | 0 | 4.41 | 0.47 | 0.77 | 2.19 | 4.45 | 0.91 | 0.98
| Analytical results for Er in OREAS 464 (Certified Value 38.2 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 36.1 | 36.7 | 39.8 | 38.8 | 37.4 | 37.8091 | 37.1 | 36.7 | 41.1742 | < 44 | 38.4 | 32.5952 | 23.2 | 36 | 38.7 | 1
| 1 | 2 | 36 | 38.2 | 40.9 | 38.8 | 38.9 | 38.5 | 37.8 | 36.7 | 43.8604 | < 44 | 37.9 | 32.8685 | 28.3 | 41 | 37.2 | 6
| 1 | 3 | 36.7 | 38.5 | 40.3 | 38.9 | 39 | 38.2374 | 37.2 | 36.4 | 42.9194 | < 44 | 36.9 | 31.904 | 28.5 | 38 | 39 | 16
| 1 | 4 | 37.4 | 38.6 | 41.7 | 40 | 38.1 | 38.0369 | 37.2 | 37.6 | 42.2179 | < 44 | 37.3 | 32.0517 | 22.5 | 38.5 | 38.8 | 38.1522
| 1 | 5 | 35.4 | 37.7 | 40.2 | 39.4 | 37.4 | 39.6188 | 37.4 | 39.8 | 45.5379 | < 44 | 38.4 | 32.509 | 29 | 38.5 | 43.8 | 28
| 1 | 6 | 36.5 | 37.5 | 40.5 | 39.4 | 39.5 | 37.505 | 37.9 | 37.3 | 45.0941 | < 44 | 38 | 32.2192 | 28.1 | 38.5 | 38.3
| Mean |  | 36.35 | 37.8667 | 40.5667 | 39.2167 | 38.3833 | 38.2845 | 37.4333 | 37.4167 | 43.4673 |  | 37.8167 | 32.3579 | 26.6 | 38.4167 | 39.3
| Median |  | 36.3 | 37.95 | 40.4 | 39.15 | 38.5 | 38.1372 | 37.3 | 37 | 43.3899 |  | 37.95 | 32.3641 | 28.2 | 38.5 | 38.75
| Std Dev. |  | 0.68337 | 0.7174 | 0.66232 | 0.47504 | 0.88412 | 0.7381 | 0.33862 | 1.2481 | 1.6857 |  | 0.6047 | 0.36299 | 2.9285 | 1.5943 | 2.2961
| Rel.Std.Dev. |  | 1.88% | 1.895% | 1.633% | 1.211% | 2.303% | 1.928% | 0.9046% | 3.336% | 3.878% |  | 1.599% | 1.122% | 11.01% | 4.15% | 5.842%
| PDM3 |  | -4.724% | -0.7485% | 6.328% | 2.79% | 0.6057% | 0.3468% | -1.884% | -1.928% | 13.93% |  | -0.8795% | -15.19% | -30.28% | 0.6931% | 3.008%
| Z-Score (Absolute) |  | 1.31 | 0.16 | 1.89 | 0.87 | 0.23 | 0.16 | 0.49 | 0.86 | 4.09 | NA | 0.2 | 4.34 | 8.7 | 0.26 | 0.25
| Analytical results for Eu in OREAS 464 (Certified Value 324 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 324 | 314 | 325 | 320 | 328 | 324.5695 | 343.8 | 317.1 | 201.2336 | 328 | 322.5 | 333.6776 | 237 | 330 | 316 | 1
| 1 | 2 | 321 | 312 | 321 | 325 | 339 | 317.6364 | 350.8 | 319.9 | 199.6557 | 328 | 326.5 | 332.7602 | 298 | 338 | 309 | 7
| 1 | 3 | 326 | 316 | 306 | 328 | 340 | 326.5925 | 343.2 | 308.6 | 199.9728 | 328 | 321.5 | 323.3296 | 301 | 328 | 300 | 16
| 1 | 4 | 331 | 347 | 329 | 325 | 332 | 313.1858 | 341.8 | 321.4 | 202.9079 | 320 | 322.5 | 326.5935 | 235 | 321 | 316 | 323.8264
| 1 | 5 | 318 | 325 | 311 | 325 | 321 | 321.2518 | 343.1 | 312.9 | 202.0314 | 328 | 328.5 | 332.8357 | 305 | 325 | 316 | 29
| 1 | 6 | 321 | 335 | 320 | 322 | 346 | 316.3762 | 347.7 | 311.9 | 201.1757 | 328 | 328 | 325.5491 | 297 | 338 | 315
| Mean |  | 323.5 | 324.8333 | 318.6667 | 324.1667 | 334.3333 | 319.9354 | 345.0667 | 315.3 | 201.1629 | 326.6667 | 324.9167 | 329.1243 | 278.8333 | 330 | 312
| Median |  | 322.5 | 320.5 | 320.5 | 325 | 335.5 | 319.4441 | 343.5 | 315 | 201.2046 | 328 | 324.5 | 329.6768 | 297.5 | 329 | 315.5
| Std Dev. |  | 4.5935 | 13.8191 | 8.641 | 2.7869 | 9.0921 | 5.1193 | 3.4477 | 4.9763 | 1.2241 | 3.266 | 3.1051 | 4.4832 | 33.3012 | 6.8993 | 6.4807
| Rel.Std.Dev. |  | 1.42% | 4.254% | 2.712% | 0.8597% | 2.719% | 1.6% | 0.9991% | 1.578% | 0.6085% | 0.9998% | 0.9557% | 1.362% | 11.94% | 2.091% | 2.077%
| PDM3 |  | -0.1008% | 0.3109% | -1.593% | 0.1051% | 3.245% | -1.202% | 6.559% | -2.633% | -37.88% | 0.8771% | 0.3367% | 1.636% | -13.89% | 1.906% | -3.652%
| Z-Score (Absolute) |  | 0.08 | 0.08 | 0.67 | 0 | 1.25 | 0.52 | 2.56 | 1.09 | 15.08 | 0.32 | 0.09 | 0.61 | 5.56 | 0.72 | 1.2
| Analytical results for Fe in OREAS 464 (Certified Value 37.24 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 35.5236 | 36.8 | 36.61 | 37.3007 | N.A. | 38.13 | 37.3077 | 22.872 | 37.41 | 38.2 | > 25 |  |  |  |  | 1
| 1 | 2 | 35.7068 | 37.41 | 37.44 | 36.4888 | N.A. | 37.99 | 37.3009 | 36.161 | 39.02 | 37.8 | > 25 |  |  |  |  | 0
| 1 | 3 | 36.1889 | 36.95 | 37.23 | 37.2636 | N.A. | 38.15 | 37.3009 | 35.882 | 39.16 | 38.7 | > 25 |  |  |  |  | 16
| 1 | 4 | 35.0125 | 37.36 | 37.04 | 36.921 | N.A. | 38.02 | 37.5858 | 26.159 | 38.11 | 38.1 | > 25 |  |  |  |  | 37.2376
| 1 | 5 | 35.0896 | 37.07 | 37.49 | 37.0523 | N.A. | 38.06 | 37.4841 | 36.301 | 38.09 | 38.8 | > 25 |  |  |  |  | 30
| 1 | 6 | 34.7811 | 37.69 | 37.35 | 36.9481 | N.A. | 38.48 | 37.3077 | 36.301 | 37.74 | 38.9 | > 25
| Mean |  | 35.3837 | 37.2133 | 37.1933 | 36.9958 |  | 38.1383 | 37.3812 | 32.2793 | 38.255 | 38.4167
| Median |  | 35.3066 | 37.215 | 37.29 | 37.0002 |  | 38.095 | 37.3077 | 36.0215 | 38.1 | 38.45
| Std Dev. |  | 0.52149 | 0.3311 | 0.32819 | 0.29421 |  | 0.17837 | 0.1234 | 6.1049 | 0.69739 | 0.4446
| Rel.Std.Dev. |  | 1.474% | 0.8897% | 0.8824% | 0.7952% |  | 0.4677% | 0.3301% | 18.91% | 1.823% | 1.157%
| PDM3 |  | -4.979% | -0.06528% | -0.119% | -0.6496% |  | 2.419% | 0.3855% | -13.32% | 2.732% | 3.166%
| Z-Score (Absolute) |  | 1.33 | 0 | 0.01 | 0.16 | NA | 0.67 | 0.12 | 0.77 | 0.76 | 0.88 | NA
| Analytical results for Ga in OREAS 464 (Certified Value 106 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 12 | 14 | 20 | 22 | 23 | [6 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*MS | PF*OES/MS | [6 empty] | 0
0
| 1 | 1 | 117 | 98.5 | 117.5 | 109.5 | 70.422 | N.A. | < 1 | 184 | 68 | [6 empty] | 1
| 1 | 2 | 118 | 102.5 | 115 | 109.5 | 71.1984 | N.A. | < 1 | 184 | 69 | [6 empty] | 21
| 1 | 3 | 117 | 100.5 | 119.5 | 107.5 | 70.6778 | N.A. | < 1 | 178 | 62 | [6 empty] | 16
| 1 | 4 | 125 | 104 | 119 | 105.5 | 70.6153 | N.A. | < 1 | 172 | 64 | [6 empty] | 105.8017
| 1 | 5 | 116.5 | 105.5 | 119 | 150.5 | 71.4563 | N.A. | < 1 | 174 | 73 | [6 empty] | 31
| 1 | 6 | 118.5 | 102 | 113.5 | 145.5 | 70.7797 | N.A. | < 1 | 180 | 69
| Mean |  | 118.6667 | 102.1667 | 117.25 | 121.3333 | 70.8583 |  |  | 178.6667 | 67.5
| Median |  | 117.5 | 102.25 | 118.25 | 109.5 | 70.7287 |  |  | 179 | 68.5
| Std Dev. |  | 3.1885 | 2.4833 | 2.4648 | 20.7694 | 0.39026 |  |  | 5.0067 | 3.937
| Rel.Std.Dev. |  | 2.687% | 2.431% | 2.102% | 17.12% | 0.5508% |  |  | 2.802% | 5.833%
| PDM3 |  | 12.16% | -3.436% | 10.82% | 14.68% | -33.03% |  |  | 68.87% | -36.2%
| Z-Score (Absolute) |  | 0.01 | 0.67 | 0 | 0.18 | 2.07 | NA | NA | 2.75 | 2.22
| Analytical results for Gd in OREAS 464 (Certified Value 676 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 680 | 681 | 663 | 621 | 690 | 665.7 | 626.54 | 712.8559 | 720 | 698.5 | 512.8 | 782 | 678 | 638 |  | 1
| 1 | 2 | 682 | 674 | 673 | 634 | 728 | 679.8 | 644.69 | 711.1402 | 720 | 700 | 530.1 | 839 | 710 | 624 |  | 8
| 1 | 3 | 680 | 689 | 663 | 636 | 680 | 666.4 | 627.5 | 713.8237 | 711 | 691 | 534.4 | 856 | 680 | 610 |  | 16
| 1 | 4 | 716 | 692 | 690 | 640 | 713 | 670.5 | 632.09 | 712.2743 | 694 | 693 | 529.2 | 861 | 692 | 633 |  | 676.0411
| 1 | 5 | 695 | 687 | 672 | 635 | 695 | 669.6 | 629.59 | 710.4292 | 711 | 705 | 542.3 | 858 | 678 | 652 |  | 32
| 1 | 6 | 678 | 671 | 674 | 629 | 659 | 671.4 | 630.19 | 715.3256 | 711 | 712 | 540.5 | 823 | 692 | 643
| Mean |  | 688.5 | 682.3333 | 672.5 | 632.5 | 694.1667 | 670.5667 | 631.7667 | 712.6415 | 711.1667 | 699.9167 | 531.55 | 836.5 | 688.3333 | 633.3333
| Median |  | 681 | 684 | 672.5 | 634.5 | 692.5 | 670.05 | 629.89 | 712.5651 | 711 | 699.25 | 532.25 | 847.5 | 686 | 635.5
| Std Dev. |  | 14.802 | 8.4774 | 9.8944 | 6.6558 | 24.3098 | 5.0583 | 6.6312 | 1.785 | 9.4956 | 7.7616 | 10.6111 | 30.3364 | 12.4847 | 14.8009
| Rel.Std.Dev. |  | 2.15% | 1.242% | 1.471% | 1.052% | 3.502% | 0.7543% | 1.05% | 0.2505% | 1.335% | 1.109% | 1.996% | 3.627% | 1.814% | 2.337%
| PDM3 |  | 1.843% | 0.9307% | -0.5238% | -6.441% | 2.681% | -0.8098% | -6.549% | 5.414% | 5.196% | 3.532% | -21.37% | 23.74% | 1.818% | -6.317%
| Z-Score (Absolute) |  | 0.01 | 0.01 | 0.3 | 1.47 | 0.34 | 0.35 | 1.49 | 0.88 | 0.84 | 0.51 | 4.43 | 4.51 | 0.17 | 1.45
| Analytical results for Ge in OREAS 464 (Indicative Value 43.3 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 20 | 22 | 23 | [12 empty] | ppm
|  |  | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 1
1
| 1 | 1 | < 1 | 40 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | 60 | < 1 | [12 empty] | 15
| 1 | 3 | < 1 | 40 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | 40 | < 1 | [12 empty] | 43.3333
| 1 | 5 | < 1 | 40 | < 1 | [12 empty] | 21
| 1 | 6 | < 1 | 40 | < 1
| Mean |  |  | 43.3333
| Median |  |  | 40
| Std Dev. |  |  | 8.165
| Rel.Std.Dev. |  |  | 18.84%
| PDM3 |  |  | 8.882e-14%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Hf in OREAS 464 (Certified Value 12.5 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 20 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 13 | 13 | 13 | 12.9 | 12.7 | 13.3 | 10.2 | N.A. | 12 | 11 | 12 |  |  |  |  | 1
| 1 | 2 | 13 | 12.3 | 13 | 12.7 | 13 | 12.8 | 10 | N.A. | 12 | 13 | 11 |  |  |  |  | 9
| 1 | 3 | 13 | 13.3 | 14 | 12.6 | 12.2 | 12.6 | 9.7 | N.A. | 11 | 13 | 12 |  |  |  |  | 16
| 1 | 4 | 13 | 14.5 | 14 | 13 | 12.3 | 12.5 | 9.7 | N.A. | 12 | 11 | 12 |  |  |  |  | 12.5167
| 1 | 5 | 12 | 12.7 | 14 | 13 | 12.1 | 12.6 | 9.9 | N.A. | 12 | 13 | 12 |  |  |  |  | 33
| 1 | 6 | 13 | 12.3 | 13 | 12.1 | 11.6 | 12.5 | 10 | N.A. | 12 | 12 | 11
| Mean |  | 12.8333 | 13.0167 | 13.5 | 12.7167 | 12.3167 | 12.7167 | 9.9167 |  | 11.8333 | 12.1667 | 11.6667
| Median |  | 13 | 12.85 | 13.5 | 12.8 | 12.25 | 12.6 | 9.95 |  | 12 | 12.5 | 12
| Std Dev. |  | 0.40825 | 0.82563 | 0.54772 | 0.34303 | 0.48751 | 0.30605 | 0.19408 |  | 0.40825 | 0.98319 | 0.5164
| Rel.Std.Dev. |  | 3.181% | 6.343% | 4.057% | 2.697% | 3.958% | 2.407% | 1.957% |  | 3.45% | 8.081% | 4.426%
| PDM3 |  | 2.53% | 3.995% | 7.856% | 1.598% | -1.598% | 1.598% | -20.77% |  | -5.459% | -2.796% | -6.791%
| Z-Score (Absolute) |  | 0.54 | 0.81 | 1.51 | 0.37 | 0.2 | 0.2 | 3.67 | NA | 0.9 | 0.42 | 1.14
| Analytical results for Ho in OREAS 464 (Certified Value 21.3 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 20.1 | 21 | 22.6 | 21.5 | 20.8 | 22.5977 | 21.7 | 20.4 | 14.8518 | < 44 | 21.5 | 18.9899 | 21.8 | 22.8 | 19.4 | 1
| 1 | 2 | 20.3 | 20.6 | 22.4 | 22 | 21.2 | 22.4039 | 21.8 | 20.7 | 14.0431 | < 44 | 21.2 | 19.1106 | 23.1 | 23.4 | 19.1 | 10
| 1 | 3 | 20.1 | 21.3 | 22.4 | 21.6 | 22.1 | 22.9464 | 22.1 | 20.9 | 14.4167 | < 44 | 21 | 18.5419 | 23.6 | 22.2 | 18.7 | 16
| 1 | 4 | 20.8 | 23.2 | 23.5 | 22 | 21.1 | 21.9271 | 21.7 | 20.9 | 14.1791 | < 44 | 21 | 18.4882 | 21.4 | 22.2 | 19.1 | 21.3049
| 1 | 5 | 19.7 | 21.7 | 22.7 | 22 | 20.3 | 22.8083 | 22 | 21.9 | 14.4088 | < 44 | 21.3 | 19.0365 | 23.6 | 22.2 | 19.4 | 34
| 1 | 6 | 20.2 | 20.3 | 23.2 | 21.9 | 21.7 | 22.2302 | 21.9 | 20.7 | 14.1305 | < 44 | 21.6 | 18.5828 | 22.5 | 22.8 | 19.4
| Mean |  | 20.2 | 21.35 | 22.8 | 21.8333 | 21.2 | 22.4856 | 21.8667 | 20.9167 | 14.3383 |  | 21.2667 | 18.7917 | 22.6667 | 22.6 | 19.1833
| Median |  | 20.15 | 21.15 | 22.65 | 21.95 | 21.15 | 22.5008 | 21.85 | 20.8 | 14.294 |  | 21.25 | 18.7864 | 22.8 | 22.5 | 19.25
| Std Dev. |  | 0.35777 | 1.033 | 0.45166 | 0.22509 | 0.63875 | 0.3776 | 0.1633 | 0.51543 | 0.29343 |  | 0.25033 | 0.28251 | 0.9288 | 0.4899 | 0.27869
| Rel.Std.Dev. |  | 1.771% | 4.838% | 1.981% | 1.031% | 3.013% | 1.679% | 0.7468% | 2.464% | 2.046% |  | 1.177% | 1.503% | 4.098% | 2.168% | 1.453%
| PDM3 |  | -5.186% | 0.2116% | 7.018% | 2.48% | -0.4925% | 5.542% | 2.637% | -1.822% | -32.7% |  | -0.1795% | -11.8% | 6.392% | 6.079% | -9.958%
| Z-Score (Absolute) |  | 0.65 | 0.02 | 0.88 | 0.31 | 0.06 | 0.69 | 0.33 | 0.35 | 4.11 | NA | 0.02 | 1.48 | 0.8 | 0.76 | 1.25
| Analytical results for In in OREAS 464 (Certified Value 2.14 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 2.0433 | 2.1 | 2 | 1.5 | 1.8 | 2.3 | [9 empty] | 1
| 1 | 2 | 2.1468 | 2.2 | 2 | 1.7 | 2.2 | 2.2 | [9 empty] | 0
| 1 | 3 | 2.5078 | 2.1 | 2 | 1.7 | 2 | 2.1 | [9 empty] | 16
| 1 | 4 | 2.2575 | 2.1 | 2 | 1.5 | 2 | 2.6 | [9 empty] | 2.1418
| 1 | 5 | 2.2685 | 2.3 | 2 | 1.7 | 2.2 | 2.2 | [9 empty] | 35
| 1 | 6 | 2.0491 | 2.1 | 2 | 1.6 | 2.2 | 3.5
| Mean |  | 2.2122 | 2.15 | 2 | 1.6167 | 2.0667 | 2.4833
| Median |  | 2.2022 | 2.1 | 2 | 1.65 | 2.1 | 2.25
| Std Dev. |  | 0.17438 | 0.083666 | 0 | 0.098319 | 0.1633 | 0.52694
| Rel.Std.Dev. |  | 7.883% | 3.891% | 0% | 6.082% | 7.902% | 21.22%
| PDM3 |  | 3.287% | 0.3844% | -6.619% | -24.52% | -3.506% | 15.95%
| Z-Score (Absolute) |  | 0.66 | 0.26 | 0.69 | 3.13 | 0.26 | 1.09
| Analytical results for K in OREAS 464 (Certified Value 0.086 wt.%) | [16 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.22 | 0.12 | 0.08 | 0.098033 | N.A. | < 0.1 | 0.07148 | 0.083 | 0.08 | < 0.1 | [5 empty] | 1
| 1 | 2 | 0.217 | 0.12 | 0.07 | 0.097925 | N.A. | < 0.1 | 0.06982 | 0.083 | 0.08 | < 0.1 | [5 empty] | 11
| 1 | 3 | 0.227 | 0.11 | 0.07 | 0.098614 | N.A. | < 0.1 | 0.07148 | 0.083 | 0.08 | < 0.1 | [5 empty] | 16
| 1 | 4 | 0.211 | 0.11 | 0.08 | 0.098723 | N.A. | < 0.1 | 0.07065 | 0.091 | 0.07 | < 0.1 | [5 empty] | 0.086243
| 1 | 5 | 0.223 | 0.11 | 0.08 | 0.097515 | N.A. | < 0.1 | 0.0733 | 0.083 | 0.07 | < 0.1 | [5 empty] | 36
| 1 | 6 | 0.25 | 0.11 | 0.08 | 0.097692 | N.A. | < 0.1 | 0.07314 | 0.083 | 0.06 | < 0.1
| Mean |  | 0.22467 | 0.11333 | 0.076667 | 0.098084 |  |  | 0.071645 | 0.084333 | 0.073333
| Median |  | 0.2215 | 0.11 | 0.08 | 0.097979 |  |  | 0.07148 | 0.083 | 0.075
| Std Dev. |  | 0.013545 | 0.005164 | 0.005164 | 0.00048881 |  |  | 0.0013674 | 0.003266 | 0.008165
| Rel.Std.Dev. |  | 6.029% | 4.556% | 6.736% | 0.4984% |  |  | 1.909% | 3.873% | 11.13%
| PDM3 |  | 160.5% | 31.41% | -11.1% | 13.73% |  |  | -16.93% | -2.214% | -14.97%
| Z-Score (Absolute) |  | 7.15 | 1.53 | 0.41 | 0.72 | NA | NA | 0.67 | 0 | 0.58 | NA
| Analytical results for La in OREAS 464 (Certified Value 1.17 wt.%) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.185 | 1.13 | 1.17 | > 1 | 1.165 | 1.077 | 1.1716 | 1.1704 | 1.1674 | 1.151 | 1.1986 | 1.2604 | 1.1 | 1.19 | 1.15 | 1
| 1 | 2 | 1.17 | 1.185 | 1.175 | > 1 | 1.175 | 1.07 | 1.2083 | 1.1795 | 1.1091 | 1.16 | 1.1993 | 1.2603 | 1.15 | 1.2 | 1.09 | 0
| 1 | 3 | 1.215 | 1.18 | 1.115 | > 1 | 1.195 | 1.1071 | 1.1921 | 1.179 | 1.0942 | 1.16 | 1.1918 | 1.2263 | 1.19 | 1.2 | 1.12 | 16
| 1 | 4 | 1.215 | 1.13 | 1.205 | > 1 | 1.18 | 1.0531 | 1.1854 | 1.1752 | 1.1725 | 1.134 | 1.192 | 1.2466 | 1.14 | 1.21 | 1.25 | 1.1706
| 1 | 5 | 1.175 | 1.14 | 1.15 | > 1 | 1.17 | 1.1017 | 1.1944 | 1.1838 | 1.1435 | 1.151 | 1.2161 | 1.2781 | 1.18 | 1.17 | 1.21 | 37
| 1 | 6 | 1.175 | 1.125 | 1.17 | > 1 | 1.23 | 1.0638 | 1.1831 | 1.1805 | 1.1463 | 1.143 | 1.2303 | 1.2609 | 1.14 | 1.2 | 1.16
| Mean |  | 1.1892 | 1.1483 | 1.1642 |  | 1.1858 | 1.0788 | 1.1892 | 1.1781 | 1.1388 | 1.1498 | 1.2047 | 1.2554 | 1.15 | 1.195 | 1.1633
| Median |  | 1.18 | 1.135 | 1.17 |  | 1.1775 | 1.0735 | 1.1888 | 1.1792 | 1.1449 | 1.151 | 1.1989 | 1.2604 | 1.145 | 1.2 | 1.155
| Std Dev. |  | 0.020595 | 0.026957 | 0.029903 |  | 0.023962 | 0.021406 | 0.012326 | 0.0046629 | 0.031307 | 0.010068 | 0.015358 | 0.017427 | 0.032249 | 0.013784 | 0.058538
| Rel.Std.Dev. |  | 1.732% | 2.347% | 2.569% |  | 2.021% | 1.984% | 1.037% | 0.3958% | 2.749% | 0.8756% | 1.275% | 1.388% | 2.804% | 1.153% | 5.032%
| PDM3 |  | 1.584% | -1.904% | -0.5512% |  | 1.3% | -7.844% | 1.585% | 0.6362% | -2.715% | -1.776% | 2.91% | 7.244% | -1.761% | 2.083% | -0.6224%
| Z-Score (Absolute) |  | 0.6 | 0.72 | 0.21 | NA | 0.21 | 2.98 | 0.6 | 0.24 | 1.03 | 0.68 | 1.11 | 2.76 | 0.67 | 0.79 | 0.24
| Analytical results for Li in OREAS 464 (Indicative Value 8.13 ppm) | [16 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 8.8475 | 8 | 9 | < 10 | 7 | < 10 | [9 empty] | 1
| 1 | 2 | 8.8174 | 8 | 7 | < 10 | 7 | < 10 | [9 empty] | 16
| 1 | 3 | 8.1222 | 8 | 8 | < 10 | 8 | < 10 | [9 empty] | 16
| 1 | 4 | 8.8315 | 8 | 8 | < 10 | 7 | < 10 | [9 empty] | 8.1337
| 1 | 5 | 9.1703 | 8 | 8 | < 10 | 8 | < 10 | [9 empty] | 22
| 1 | 6 | 8.6747 | 8 | 9 | < 10 | 8 | < 10
| Mean |  | 8.7439 | 8 | 8.1667 |  | 7.5
| Median |  | 8.8245 | 8 | 8 |  | 7.5
| Std Dev. |  | 0.34554 | 0 | 0.75277 |  | 0.54772
| Rel.Std.Dev. |  | 3.952% | 0% | 9.218% |  | 7.303%
| PDM3 |  | 7.502% | -1.644% | 0.4048% |  | -7.791%
| Z-Score (Absolute) |  | 1.59 | 0.17 | 0.17 | NA | 1.18 | NA
| Analytical results for Lu in OREAS 464 (Certified Value 1.69 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 1.8 | 1.67 | 1.61 | 1.76 | 1.71 | 1.7881 | 1.6 | 1.7 | < 44 | 2 | 1.4486 | 0.89 | 1.6 | 1.92 |  | 1
| 1 | 2 | 1.73 | 1.69 | 1.74 | 1.78 | 1.72 | 1.9302 | 1.5 | 1.5 | < 44 | 2 | 1.4469 | 0.92 | 1.6 | 1.95 |  | 1
| 1 | 3 | 1.75 | 1.58 | 1.62 | 1.76 | 1.62 | 1.6976 | 1.7 | 1.6 | < 44 | 2 | 1.4317 | 0.95 | 1.6 | 1.96 |  | 16
| 1 | 4 | 1.79 | 1.64 | 1.75 | 1.78 | 1.73 | 1.5037 | 1.6 | 1.6 | < 44 | 2 | 1.3902 | 0.77 | 1.8 | 2 |  | 1.6859
| 1 | 5 | 1.67 | 1.66 | 1.59 | 1.84 | 1.74 | 1.8284 | 1.7 | 1.9 | < 44 | 2 | 1.4463 | 0.89 | 1.6 | 2.2 |  | 38
| 1 | 6 | 1.69 | 1.61 | 1.72 | 1.87 | 1.65 | 1.6233 | 1.6 | 1.6 | < 44 | 1.5 | 1.3974 | 0.92 | 1.6 | 1.92
| Mean |  | 1.7383 | 1.6417 | 1.6717 | 1.7983 | 1.695 | 1.7285 | 1.6167 | 1.65 |  | 1.9167 | 1.4269 | 0.89 | 1.6333 | 1.9917
| Median |  | 1.74 | 1.65 | 1.67 | 1.78 | 1.715 | 1.7429 | 1.6 | 1.6 |  | 2 | 1.439 | 0.905 | 1.6 | 1.955
| Std Dev. |  | 0.052313 | 0.040702 | 0.072503 | 0.045789 | 0.048477 | 0.15282 | 0.075277 | 0.13784 |  | 0.20412 | 0.026418 | 0.062929 | 0.08165 | 0.10629
| Rel.Std.Dev. |  | 3.009% | 2.479% | 4.337% | 2.546% | 2.86% | 8.841% | 4.656% | 8.354% |  | 10.65% | 1.852% | 7.071% | 4.999% | 5.337%
| PDM3 |  | 3.107% | -2.627% | -0.8471% | 6.666% | 0.5369% | 2.526% | -4.109% | -2.132% |  | 13.68% | -15.37% | -47.21% | -3.121% | 18.13%
| Z-Score (Absolute) |  | 0.79 | 0.36 | 0 | 1.5 | 0.28 | 0.67 | 0.65 | 0.26 | NA | 3.89 | 2.9 | 9.27 | 0.45 | 3.79
| Analytical results for Mg in OREAS 464 (Certified Value 0.846 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.85376 | 0.87 | 0.87 | 0.80939 | N.A. | 0.89 | 0.77856 | 0.766 | 0.88 | 0.86 | 0.85 |  |  |  |  | 1
| 1 | 2 | 0.84653 | 0.88 | 0.88 | 0.79582 | N.A. | 0.885 | 0.77313 | 0.802 | 0.86 | 0.85 | 0.82 |  |  |  |  | 2
| 1 | 3 | 0.861 | 0.89 | 0.88 | 0.80693 | N.A. | 0.89 | 0.77192 | 0.802 | 0.87 | 0.9 | 0.85 |  |  |  |  | 16
| 1 | 4 | 0.84239 | 0.88 | 0.88 | 0.80897 | N.A. | 0.89 | 0.78519 | 0.784 | 0.86 | 0.86 | 0.84 |  |  |  |  | 0.84555
| 1 | 5 | 0.84446 | 0.87 | 0.88 | 0.80561 | N.A. | 0.895 | 0.77494 | 0.802 | 0.85 | 0.88 | 0.85 |  |  |  |  | 39
| 1 | 6 | 0.83413 | 0.87 | 0.88 | 0.80656 | N.A. | 0.895 | 0.78278 | 0.808 | 0.84 | 0.88 | 0.85
| Mean |  | 0.84705 | 0.87667 | 0.87833 | 0.80555 |  | 0.89083 | 0.77775 | 0.794 | 0.86 | 0.87167 | 0.84333
| Median |  | 0.8455 | 0.875 | 0.88 | 0.80675 |  | 0.89 | 0.77675 | 0.802 | 0.86 | 0.87 | 0.85
| Std Dev. |  | 0.0093312 | 0.008165 | 0.0040825 | 0.0049839 |  | 0.0037639 | 0.0053766 | 0.01595 | 0.014142 | 0.018348 | 0.012111
| Rel.Std.Dev. |  | 1.102% | 0.9314% | 0.4648% | 0.6187% |  | 0.4225% | 0.6913% | 2.009% | 1.644% | 2.105% | 1.436%
| PDM3 |  | 0.1769% | 3.68% | 3.877% | -4.731% |  | 5.355% | -8.018% | -6.097% | 1.709% | 3.089% | -0.2622%
| Z-Score (Absolute) |  | 0.2 | 0.65 | 0.7 | 1.39 | NA | 1.06 | 2.19 | 1.56 | 0.17 | 0.51 | 0.17
| Analytical results for Mn in OREAS 464 (Certified Value 0.282 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.3 | 0.28 | 0.29107 | N.A. | 0.31 | 0.25922 | 0.178 | 0.312 | 0.284 | 0.256 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.28 | 0.28848 | N.A. | 0.305 | 0.25534 | 0.271 | 0.302 | 0.28 | 0.249 | [5 empty] | 3
| 1 | 3 | 0.3 | 0.28 | 0.27832 | N.A. | 0.31 | 0.25697 | 0.263 | 0.308 | 0.288 | 0.245 | [5 empty] | 16
| 1 | 4 | 0.3 | 0.28 | 0.27787 | N.A. | 0.31 | 0.25929 | 0.201 | 0.298 | 0.284 | 0.253 | [5 empty] | 0.28243
| 1 | 5 | 0.3 | 0.28 | 0.28805 | N.A. | 0.31 | 0.25929 | 0.271 | 0.294 | 0.29 | 0.253 | [5 empty] | 40
| 1 | 6 | 0.3 | 0.28 | 0.28798 | N.A. | 0.31 | 0.25968 | 0.271 | 0.296 | 0.29 | 0.258
| Mean |  | 0.3 | 0.28 | 0.2853 |  | 0.30917 | 0.2583 | 0.2425 | 0.30167 | 0.286 | 0.25233
| Median |  | 0.3 | 0.28 | 0.28802 |  | 0.31 | 0.25926 | 0.267 | 0.3 | 0.286 | 0.253
| Std Dev. |  | 0 | 0 | 0.0056941 |  | 0.0020412 | 0.0017459 | 0.041808 | 0.0070899 | 0.004 | 0.0047188
| Rel.Std.Dev. |  | 0% | 0% | 1.996% |  | 0.6602% | 0.6759% | 17.24% | 2.35% | 1.399% | 1.87%
| PDM3 |  | 6.22% | -0.8612% | 1.014% |  | 9.466% | -8.545% | -14.14% | 6.81% | 1.263% | -10.66%
| Z-Score (Absolute) |  | 0.61 | 0.22 | 0 | NA | 1 | 1.13 | 0.67 | 0.68 | 0.03 | 1.37
| Analytical results for Mo in OREAS 464 (Certified Value 77 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
1
| 1 | 1 | 78 | 37.156 | N.A. | 80 | 72.7564 | 70 | 75 | 84 | [7 empty] | 1
| 1 | 2 | 76 | 37.2795 | N.A. | 80 | 74.9987 | 72 | 75 | 81 | [7 empty] | 0
| 1 | 3 | 77 | 36.5775 | N.A. | 80 | 71.7111 | 74 | 70 | 82 | [7 empty] | 16
| 1 | 4 | 78 | 37.8922 | N.A. | 85 | 72.88 | 76 | 75 | 84 | [7 empty] | 77.0732
| 1 | 5 | 77 | 36.3622 | N.A. | 80 | 75.1085 | 74 | 75 | 81 | [7 empty] | 41
| 1 | 6 | 76 | 37.841 | N.A. | 85 | 75.1795 | 70 | 80 | 82
| Mean |  | 77 | 37.1848 |  | 81.6667 | 73.7724 | 72.6667 | 75 | 82.3333
| Median |  | 77 | 37.2178 |  | 80 | 73.9393 | 73 | 75 | 82
| Std Dev. |  | 0.89443 | 0.63028 |  | 2.582 | 1.5064 | 2.4221 | 3.1623 | 1.3663
| Rel.Std.Dev. |  | 1.162% | 1.695% |  | 3.162% | 2.042% | 3.333% | 4.216% | 1.659%
| PDM3 |  | -0.09494% | -51.75% |  | 5.96% | -4.283% | -5.717% | -2.69% | 6.825%
| Z-Score (Absolute) |  | 0.58 | 10.93 | NA | 1.93 | 0.35 | 0.67 | 0 | 2.12
| Analytical results for Na in OREAS 464 (Indicative Value 0.135 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 11 | 14 | 18 | 20 | [11 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES | BF*OES | [11 empty] | 3
3
| 1 | 1 | 0.14 | N.A. | 0.12864 | 0.69 | [11 empty] | 1
| 1 | 2 | 0.14 | N.A. | 0.12856 | 0.67 | [11 empty] | 4
| 1 | 3 | 0.14 | N.A. | 0.1296 | 0.67 | [11 empty] | 16
| 1 | 4 | 0.15 | N.A. | 0.12908 | 0.66 | [11 empty] | 0.13501
| 1 | 5 | 0.14 | N.A. | 0.1293 | 0.63 | [11 empty] | 23
| 1 | 6 | 0.14 | N.A. | 0.12493 | 0.63
| Mean |  | 0.14167 |  | 0.12835 | 0.65833
| Median |  | 0.14 |  | 0.12886 | 0.665
| Std Dev. |  | 0.0040825 |  | 0.0017235 | 0.024014
| Rel.Std.Dev. |  | 2.882% |  | 1.343% | 3.648%
| PDM3 |  | 4.932% |  | -4.932% | 387.6%
| Z-Score (Absolute) |  | 0 | NA | 0.67 | 26.16
| Analytical results for Nb in OREAS 464 (Certified Value 1903 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 1920 | 1850 | 1850 | 1925 | 1815 | 1732 | 1800 | N.A. | 1540 | 1382.7 | 983 | 1890 | 2050 |  |  | 1
| 1 | 2 | 1865 | 1860 | 1850 | 1945 | 1875 | 1920 | 1831 | N.A. | 1630 | 1370.8 | 1460 | 1900 | 2050 |  |  | 5
| 1 | 3 | 2000 | 1940 | 1830 | 1975 | 1790 | 1897 | 1816 | N.A. | 1580 | 1265.3 | 1490 | 1950 | 2020 |  |  | 16
| 1 | 4 | 1995 | 1820 | 1900 | 1970 | 1820 | 1956 | 1802 | N.A. | 1525 | 1389.7 | 918 | 1910 | 2040 |  |  | 1903.4074
| 1 | 5 | 1930 | 1830 | 1870 | 1985 | 1760 | 1955 | 1810 | N.A. | 1635 | 1530.9 | 1480 | 1920 | 2290 |  |  | 42
| 1 | 6 | 1910 | 1870 | 1890 | 1925 | 1705 | 1942 | 1823 | N.A. | 1670 | 1247.8 | 1430 | 1970 | 2080
| Mean |  | 1936.6667 | 1861.6667 | 1865 | 1954.1667 | 1794.1667 | 1900.3333 | 1813.6667 |  | 1596.6667 | 1364.5333 | 1293.5 | 1923.3333 | 2088.3333
| Median |  | 1925 | 1855 | 1860 | 1957.5 | 1802.5 | 1931 | 1813 |  | 1605 | 1376.75 | 1445 | 1915 | 2050
| Std Dev. |  | 52.1217 | 42.6224 | 26.6458 | 26.1566 | 57.9152 | 85.5071 | 12.0775 |  | 57.5905 | 102.1317 | 267.2675 | 30.7679 | 100.681
| Rel.Std.Dev. |  | 2.691% | 2.289% | 1.429% | 1.339% | 3.228% | 4.5% | 0.6659% |  | 3.607% | 7.485% | 20.66% | 1.6% | 4.821%
| PDM3 |  | 1.747% | -2.193% | -2.018% | 2.667% | -5.739% | -0.1615% | -4.715% |  | -16.12% | -28.31% | -32.04% | 1.047% | 9.716%
| Z-Score (Absolute) |  | 0.69 | 0.02 | 0.02 | 0.85 | 0.65 | 0.66 | 0.47 | NA | 2.5 | 4.67 | 5.34 | 0.56 | 1.73
| Analytical results for Nd in OREAS 464 (Certified Value 0.994 wt.%) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 3
3
| 1 | 1 | 0.992 | 0.968 | > 1 | 0.967 | 0.981 | 1.0013 | 0.97447 | 0.95103 | 0.7653 | 1.012 | 0.982 | 0.89356 | 0.91 | 1.04 | 1.04 | 1
| 1 | 2 | 0.984 | 0.969 | > 1 | 0.973 | 1 | 1.01 | 0.9889 | 0.96106 | 0.76654 | 1.012 | 0.9916 | 0.90634 | 0.966 | 1.06 | 1.03 | 0
| 1 | 3 | 1.01 | > 1 | 0.994 | 0.984 | 1.03 | 1.0388 | 0.99765 | 0.93281 | 0.76941 | 1.02 | 0.9764 | 0.90028 | 0.996 | 1.02 | 1.01 | 16
| 1 | 4 | 1.015 | > 1 | > 1 | 0.98 | 1.015 | 1.0087 | 0.9737 | 0.95524 | 0.76679 | 0.995 | 0.9791 | 0.8974 | 0.985 | 1.03 | 1.05 | 0.99395
| 1 | 5 | 0.986 | 0.998 | > 1 | 0.986 | 0.964 | 1.0246 | 0.98859 | 0.94483 | 0.77217 | 1.012 | 0.9984 | 0.91517 | 0.989 | 1.02 | 1.06 | 43
| 1 | 6 | 0.993 | 0.924 | > 1 | 0.977 | 1.025 | 0.99982 | 0.99934 | 0.93494 | 0.76379 | 0.995 | 0.9994 | 0.91247 | 0.951 | 1.04 | 1.06
| Mean |  | 0.99667 | 0.96475 | 0.994 | 0.97783 | 1.0025 | 1.0139 | 0.98711 | 0.94665 | 0.76733 | 1.0077 | 0.98782 | 0.9042 | 0.96617 | 1.035 | 1.0417
| Median |  | 0.9925 | 0.9685 | 0.994 | 0.9785 | 1.0075 | 1.0093 | 0.98874 | 0.94793 | 0.76667 | 1.012 | 0.9868 | 0.90331 | 0.9755 | 1.035 | 1.045
| Std Dev. |  | 0.012832 | 0.030522 |  | 0.0070828 | 0.02596 | 0.015063 | 0.011006 | 0.011248 | 0.0030124 | 0.010289 | 0.010006 | 0.0085819 | 0.032096 | 0.015166 | 0.019408
| Rel.Std.Dev. |  | 1.288% | 3.164% |  | 0.7243% | 2.589% | 1.486% | 1.115% | 1.188% | 0.3926% | 1.021% | 1.013% | 0.9491% | 3.322% | 1.465% | 1.863%
| PDM3 |  | 0.273% | -2.938% | 0.004721% | -1.622% | 0.8599% | 2.002% | -0.6886% | -4.759% | -22.8% | 1.38% | -0.6174% | -9.03% | -2.796% | 4.13% | 4.8%
| Z-Score (Absolute) |  | 0.28 | 0.72 | 0.19 | 0.31 | 0.46 | 0.81 | 0.02 | 1.28 | 6.87 | 0.61 | 0 | 2.6 | 0.67 | 1.47 | 1.68
| Analytical results for Ni in OREAS 464 (Certified Value 82 ppm) | [16 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 53 | N.A. | 100 | 89.2644 | 51 | 80 | 100 | 76 | [7 empty] | 1
| 1 | 2 | 90 | N.A. | < 100 | 95.8059 | 74 | 80 | 150 | 77 | [7 empty] | 6
| 1 | 3 | 50 | N.A. | < 100 | 85.5476 | 76 | 80 | 100 | 69 | [7 empty] | 16
| 1 | 4 | 41 | N.A. | < 100 | 89.6346 | 59 | 80 | 100 | 71 | [7 empty] | 81.9869
| 1 | 5 | 62 | N.A. | < 100 | 89.4346 | 76 | 80 | 100 | 73 | [7 empty] | 44
| 1 | 6 | 49 | N.A. | < 100 | 91.1625 | 73 | 80 | 100 | 82
| Mean |  | 57.5 |  | 100 | 90.1416 | 68.1667 | 80 | 108.3333 | 74.6667
| Median |  | 51.5 |  | 100 | 89.5346 | 73.5 | 80 | 100 | 74.5
| Std Dev. |  | 17.3061 |  |  | 3.3394 | 10.572 | 1.5567e-14 | 20.4124 | 4.6762
| Rel.Std.Dev. |  | 30.1% |  |  | 3.705% | 15.51% | 1.946e-14% | 18.84% | 6.263%
| PDM3 |  | -29.87% |  | 21.97% | 9.946% | -16.86% | -2.423% | 32.13% | -8.929%
| Z-Score (Absolute) |  | 1.5 | NA | 1.33 | 0.67 | 0.56 | 0 | 1.33 | 0.35
| Analytical results for P in OREAS 464 (Certified Value 1.41 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 1.445 | 1.4 | 1.32 | 1.446 | N.A. | 1.41 | 1.4283 | 1.25 | 1.42 | 1.36 | [5 empty] | 1
| 1 | 2 | 1.432 | 1.42 | 1.36 | 1.4285 | N.A. | 1.425 | 1.4379 | 1.27 | 1.46 | 1.37 | [5 empty] | 7
| 1 | 3 | 1.44 | 1.39 | 1.37 | 1.3817 | N.A. | 1.5 | 1.4357 | 1.22 | 1.5 | 1.32 | [5 empty] | 16
| 1 | 4 | 1.406 | 1.4 | 1.34 | 1.3783 | N.A. | 1.395 | 1.4243 | 1.26 | 1.5 | 1.37 | [5 empty] | 1.4082
| 1 | 5 | 1.438 | 1.4 | 1.33 | 1.4264 | N.A. | 1.405 | 1.4462 | 1.22 | 1.5 | 1.36 | [5 empty] | 45
| 1 | 6 | 1.444 | 1.37 | 1.35 | 1.4187 | N.A. | 1.44 | 1.4379 | 1.24 | 1.46 | 1.34
| Mean |  | 1.4342 | 1.3967 | 1.345 | 1.4133 |  | 1.4292 | 1.435 | 1.2433 | 1.4733 | 1.3533
| Median |  | 1.439 | 1.4 | 1.345 | 1.4226 |  | 1.4175 | 1.4368 | 1.245 | 1.48 | 1.36
| Std Dev. |  | 0.014566 | 0.01633 | 0.018708 | 0.027275 |  | 0.038134 | 0.0077514 | 0.020656 | 0.03266 | 0.019664
| Rel.Std.Dev. |  | 1.016% | 1.169% | 1.391% | 1.93% |  | 2.668% | 0.5402% | 1.661% | 2.217% | 1.453%
| PDM3 |  | 1.842% | -0.8209% | -4.49% | 0.3593% |  | 1.487% | 1.903% | -11.71% | 4.623% | -3.898%
| Z-Score (Absolute) |  | 0.65 | 0.52 | 2.12 | 0 | NA | 0.05 | 0.67 | 5.27 | 1.86 | 1.86
| Analytical results for Pb in OREAS 464 (Certified Value 174 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 176 | 167 | 300 | 164.5319 | 89 | < 300 | 170 | 186 | [7 empty] | 1
| 1 | 2 | 173 | 194 | 300 | 164.7759 | 88 | < 300 | 160 | 213 | [7 empty] | 8
| 1 | 3 | 169 | 187 | 300 | 163.4401 | 88 | < 300 | 160 | 174 | [7 empty] | 16
| 1 | 4 | 168 | 203 | 200 | 163.5161 | 93 | < 300 | 160 | 176 | [7 empty] | 173.5236
| 1 | 5 | 175 | 202 | 300 | 165.1617 | 90 | < 300 | 160 | 188 | [7 empty] | 46
| 1 | 6 | 172 | 179 | 300 | 164.4826 | 82 | < 300 | 160 | 180
| Mean |  | 172.1667 | 188.6667 | 283.3333 | 164.318 | 88.3333 |  | 161.6667 | 186.1667
| Median |  | 172.5 | 190.5 | 300 | 164.5072 | 88.5 |  | 160 | 183
| Std Dev. |  | 3.1885 | 13.9809 | 40.8248 | 0.69399 | 3.6148 |  | 4.0825 | 14.2326
| Rel.Std.Dev. |  | 1.852% | 7.41% | 14.41% | 0.4223% | 4.092% |  | 2.525% | 7.645%
| PDM3 |  | -0.782% | 8.727% | 63.28% | -5.305% | -49.09% |  | -6.833% | 7.286%
| Z-Score (Absolute) |  | 0 | 1.06 | 7.14 | 0.5 | 5.38 | NA | 0.67 | 0.55
| Analytical results for Pr in OREAS 464 (Certified Value 2597 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 2620 | 2480 | 2580 | > 1000 | 2520 | 2662.1856 | 2543.4 | 2730.9 | 2057.5729 | 2599 | 2557.5 | 2813.1966 | 2370 | 2600 | 2420 | 1
| 1 | 2 | 2600 | 2490 | 2580 | > 1000 | 2540 | 2668.379 | 2656.7 | 2721.2 | 1963.9981 | 2582 | 2588.5 | 2813.0115 | 2510 | 2630 | 2380 | 0
| 1 | 3 | 2660 | 2610 | 2460 | > 1000 | 2570 | 2732.8798 | 2612.6 | 2738 | 2135.6463 | 2607 | 2546.5 | 2772.1669 | 2610 | 2560 | 2340 | 16
| 1 | 4 | 2680 | 2490 | 2680 | > 1000 | 2530 | 2639.0016 | 2606.5 | 2748.1 | 2049.5699 | 2541 | 2547 | 2779.994 | 2590 | 2610 | 2410 | 2597.1241
| 1 | 5 | 2600 | 2480 | 2580 | > 1000 | 2460 | 2715.8963 | 2613.2 | 2727 | 2007.2758 | 2582 | 2598 | 2829.3427 | 2560 | 2610 | 2470 | 47
| 1 | 6 | 2630 | 2500 | 2600 | > 1000 | 2600 | 2641.2812 | 2624.4 | 2705.2 | 1996.9007 | 2566 | 2625.5 | 2818.5753 | 2490 | 2570 | 2430
| Mean |  | 2631.6667 | 2508.3333 | 2580 |  | 2536.6667 | 2676.6039 | 2609.4667 | 2728.4 | 2035.1606 | 2579.5 | 2577.1667 | 2804.3812 | 2521.6667 | 2596.6667 | 2408.3333
| Median |  | 2625 | 2490 | 2580 |  | 2535 | 2665.2823 | 2612.9 | 2728.95 | 2028.4228 | 2582 | 2573 | 2813.104 | 2535 | 2605 | 2415
| Std Dev. |  | 32.5064 | 50.3653 | 70.4273 |  | 47.6095 | 39.1102 | 37.0172 | 14.687 | 60.1877 | 23.7213 | 32.0494 | 22.844 | 87.2735 | 26.5832 | 44.4597
| Rel.Std.Dev. |  | 1.235% | 2.008% | 2.73% |  | 1.877% | 1.461% | 1.419% | 0.5383% | 2.957% | 0.9196% | 1.244% | 0.8146% | 3.461% | 1.024% | 1.846%
| PDM3 |  | 1.33% | -3.419% | -0.6593% |  | -2.328% | 3.06% | 0.4752% | 5.055% | -21.64% | -0.6786% | -0.7684% | 7.98% | -2.905% | -0.01761% | -7.269%
| Z-Score (Absolute) |  | 0.5 | 1.15 | 0.18 | NA | 0.59 | 1.01 | 0.24 | 1.6 | 6.32 | 0.1 | 0.13 | 2.47 | 0.76 | 0.1 | 2.06
| Analytical results for Rb in OREAS 464 (Certified Value 4.39 ppm) | [16 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  | 2
2
| 1 | 1 | 3.5 | 4 | 4.2 | 4.2 | 4 | 4.6974 | 3.5 | 4.7 | 6.9576 | N.A. | 4 | 3.3 | 5.5 | 5.2 |  | 1
| 1 | 2 | 4.5 | 4.3 | 4.4 | 4.3 | 4 | 4.7745 | 3.8 | 5.3 | 6.0712 | N.A. | 4 | 3.4 | 5.5 | 5.2 |  | 0
| 1 | 3 | 4.8 | 4.1 | 4.4 | 4.1 | 4 | 4.2735 | 4.5 | 4.7 | 5.8513 | N.A. | 4 | 3.4 | 5.5 | 3.9 |  | 16
| 1 | 4 | 5.2 | 4.6 | 4.4 | 4.2 | 4 | 4.184 | 3.8 | 4.7 | 5.9953 | N.A. | 5 | 3.3 | 5.5 | 4.6 |  | 4.3867
| 1 | 5 | 6.3 | 4.6 | 4.4 | 4 | 4 | 4.7381 | 4.4 | 5.5 | 6.1281 | N.A. | 5 | 3.4 | 5.5 | 4.4 |  | 48
| 1 | 6 | 4.5 | 3.8 | 4.4 | 3.9 | 5 | 4.3331 | 3.8 | 5.1 | 5.9305 | N.A. | 4 | 3.2 | 5.5 | 4.8
| Mean |  | 4.8 | 4.2333 | 4.3667 | 4.1167 | 4.1667 | 4.5001 | 3.9667 | 5 | 6.1557 |  | 4.3333 | 3.3333 | 5.5 | 4.6833
| Median |  | 4.65 | 4.2 | 4.4 | 4.15 | 4 | 4.5152 | 3.8 | 4.9 | 6.0333 |  | 4 | 3.35 | 5.5 | 4.7
| Std Dev. |  | 0.9252 | 0.3266 | 0.08165 | 0.1472 | 0.40825 | 0.26456 | 0.39328 | 0.35214 | 0.40498 |  | 0.5164 | 0.08165 | 0 | 0.49967
| Rel.Std.Dev. |  | 19.28% | 7.715% | 1.87% | 3.576% | 9.798% | 5.879% | 9.915% | 7.043% | 6.579% |  | 11.92% | 2.449% | 0% | 10.67%
| PDM3 |  | 9.422% | -3.496% | -0.4561% | -6.155% | -5.015% | 2.586% | -9.575% | 13.98% | 40.33% |  | -1.216% | -24.01% | 25.38% | 6.763%
| Z-Score (Absolute) |  | 0.36 | 0.36 | 0 | 0.67 | 0.54 | 0.36 | 1.08 | 1.71 | 4.39 | NA | 0.09 | 2.79 | 3.06 | 0.85
| Analytical results for Re in OREAS 464 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 17 | 22 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | [12 empty] | 1
| 1 | 2 | < 0.1 | < 0.1 | < 1 | [12 empty] | 18
| 1 | 3 | 0.2 | < 0.1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | [12 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | [12 empty] | 24
| 1 | 6 | < 0.1 | < 0.1 | < 1
| Mean |  | 0.2
| Median |  | 0.2
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 464 (Certified Value 1028 ppm) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 0
0
| 1 | 1 | 980 | 1000 | 1000 | 1000 | 1000 | [10 empty] | 1
| 1 | 2 | 940 | 1000 | 1100 | 900 | 1200 | [10 empty] | 11
| 1 | 3 | 950 | 1000 | 1000 | 900 | 1200 | [10 empty] | 16
| 1 | 4 | 980 | 1100 | 1000 | 1100 | 1000 | [10 empty] | 1028.3333
| 1 | 5 | 970 | 1100 | 1100 | 1000 | 1200 | [10 empty] | 49
| 1 | 6 | 930 | 1000 | 1100 | 900 | 1200
| Mean |  | 958.3333 | 1033.3333 | 1050 | 966.6667 | 1133.3333
| Median |  | 960 | 1000 | 1050 | 950 | 1200
| Std Dev. |  | 21.3698 | 51.6398 | 54.7723 | 81.6497 | 103.2796
| Rel.Std.Dev. |  | 2.23% | 4.997% | 5.216% | 8.447% | 9.113%
| PDM3 |  | -6.807% | 0.4862% | 2.107% | -5.997% | 10.21%
| Z-Score (Absolute) |  | 0.76 | 0 | 0.17 | 0.67 | 1.01
| Analytical results for Sb in OREAS 464 (Certified Value 1.8 ppm) | [16 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.8 | 2.9 | 2 | 1.623 | 2 | 1.4 | [9 empty] | 1
| 1 | 2 | 2 | 2.7 | 2 | 1.6574 | 2 | 1.5 | [9 empty] | 0
| 1 | 3 | 1.8 | 2.8 | 2 | 1.4668 | 2 | 1.5 | [9 empty] | 16
| 1 | 4 | 2 | 2.8 | 2 | 1.3877 | 2 | 1.5 | [9 empty] | 1.8015
| 1 | 5 | 1.9 | 3 | 2 | 1.4881 | 2 | 1.7 | [9 empty] | 50
| 1 | 6 | 2.2 | 2.7 | 2 | 1.4204 | 2 | 1.7
| Mean |  | 1.95 | 2.8167 | 2 | 1.5073 | 2 | 1.55
| Median |  | 1.95 | 2.8 | 2 | 1.4775 | 2 | 1.5
| Std Dev. |  | 0.15166 | 0.1169 | 0 | 0.10935 | 0 | 0.12247
| Rel.Std.Dev. |  | 7.777% | 4.15% | 0% | 7.255% | 0% | 7.902%
| PDM3 |  | 8.246% | 56.36% | 11.02% | -16.33% | 11.02% | -13.96%
| Z-Score (Absolute) |  | 0.07 | 2.52 | 0.07 | 1.4 | 0.07 | 1.27
| Analytical results for Sc in OREAS 464 (Indicative Value 152 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | [11 empty] | 0
0
| 1 | 1 | 150 | 150 | 140 | 170 | [11 empty] | 1
| 1 | 2 | 152 | 150 | 140 | 160 | [11 empty] | 19
| 1 | 3 | 150 | 150 | 140 | 160 | [11 empty] | 16
| 1 | 4 | 151 | 150 | 140 | 170 | [11 empty] | 151.875
| 1 | 5 | 151 | 150 | 140 | 170 | [11 empty] | 25
| 1 | 6 | 151 | 150 | 140 | 170
| Mean |  | 150.8333 | 150 | 140 | 166.6667
| Median |  | 151 | 150 | 140 | 170
| Std Dev. |  | 0.75277 | 0 | 0 | 5.164
| Rel.Std.Dev. |  | 0.4991% | 0% | 0% | 3.098%
| PDM3 |  | -0.6859% | -1.235% | -7.819% | 9.739%
| Z-Score (Absolute) |  | 0.05 | 0.05 | 1.3 | 2.02
| Analytical results for Se in OREAS 464 (Indicative Value 40 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 1
1
| 1 | 1 | < 20 | 40 | [13 empty] | 1
| 1 | 2 | < 20 | 40 | [13 empty] | 15
| 1 | 3 | < 20 | 40 | [13 empty] | 16
| 1 | 4 | < 20 | 40 | [13 empty] | 40
| 1 | 5 | < 20 | 40 | [13 empty] | 21
| 1 | 6 | < 20 | 40
| Mean |  |  | 40
| Median |  |  | 40
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Si in OREAS 464 (Certified Value 9.52 wt.%) | [16 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 9.7588 | 9.16 | 9.6978 | N.A. | 9.63 | 9.1134 | 12.575 | 9.24 | [7 empty] | 1
| 1 | 2 | 9.7449 | 9.44 | 9.6994 | N.A. | 9.56 | 9.1834 | 13.51 | 9.15 | [7 empty] | 9
| 1 | 3 | 10.2928 | 9.31 | 9.6729 | N.A. | 9.5 | 9.1684 | 13.276 | 9.33 | [7 empty] | 16
| 1 | 4 | 9.9763 | 9.35 | 9.7887 | N.A. | 9.6 | 9.2535 | 12.855 | 9.24 | [7 empty] | 9.5193
| 1 | 5 | 10.0653 | 9.47 | 9.8244 | N.A. | 9.61 | 9.2435 | 12.996 | 9.33 | [7 empty] | 51
| 1 | 6 | 9.9664 | 9.39 | 9.6981 | N.A. | 9.69 | 9.2184 | 13.603 | 9.33
| Mean |  | 9.9674 | 9.3533 | 9.7302 |  | 9.5983 | 9.1968 | 13.1358 | 9.27
| Median |  | 9.9714 | 9.37 | 9.6988 |  | 9.605 | 9.2009 | 13.136 | 9.285
| Std Dev. |  | 0.20423 | 0.11112 | 0.060998 |  | 0.064317 | 0.052523 | 0.39771 | 0.073485
| Rel.Std.Dev. |  | 2.049% | 1.188% | 0.6269% |  | 0.6701% | 0.5711% | 3.028% | 0.7927%
| PDM3 |  | 4.707% | -1.744% | 2.215% |  | 0.8297% | -3.389% | 37.99% | -2.619%
| Z-Score (Absolute) |  | 0.76 | 0.5 | 0.27 | NA | 0 | 0.82 | 7.26 | 0.67
| Analytical results for Sm in OREAS 464 (Certified Value 1498 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1490 | 1480 | 1565 | > 1000 | 1495 | 1522.4595 | 1473.9 | 1523.6 | 1299.228 | 1470 | 1495.5 | 1527.5753 | 1400 | 1450 | 1490 | 1
| 1 | 2 | 1465 | 1515 | 1550 | > 1000 | 1515 | 1517.8858 | 1521.1 | 1550 | 1228.8764 | 1470 | 1507 | 1500.8182 | 1490 | 1500 | 1470 | 0
| 1 | 3 | 1495 | 1535 | 1480 | > 1000 | 1570 | 1556.1504 | 1487.1 | 1500.9 | 1206.7578 | 1470 | 1486 | 1484.9965 | 1540 | 1460 | 1430 | 16
| 1 | 4 | 1515 | 1450 | 1605 | > 1000 | 1535 | 1487.6657 | 1481.9 | 1543.5 | 1313.4883 | 1470 | 1482 | 1490.9371 | 1390 | 1450 | 1490 | 1498.3626
| 1 | 5 | 1460 | 1475 | 1550 | > 1000 | 1470 | 1528.289 | 1488.2 | 1542.8 | 1285.7261 | 1470 | 1520.5 | 1505.2964 | 1510 | 1450 | 1510 | 52
| 1 | 6 | 1470 | 1475 | 1545 | > 1000 | 1550 | 1532.8391 | 1509.9 | 1535.3 | 1283.8582 | 1470 | 1531 | 1476.0362 | 1460 | 1500 | 1520
| Mean |  | 1482.5 | 1488.3333 | 1549.1667 |  | 1522.5 | 1524.2149 | 1493.6833 | 1532.6833 | 1269.6558 | 1470 | 1503.6667 | 1497.6099 | 1465 | 1468.3333 | 1485
| Median |  | 1480 | 1477.5 | 1550 |  | 1525 | 1525.3742 | 1487.65 | 1539.05 | 1284.7921 | 1470 | 1501.25 | 1495.8777 | 1475 | 1455 | 1490
| Std Dev. |  | 21.1542 | 30.93 | 40.4248 |  | 36.7083 | 22.3151 | 17.9905 | 17.993 | 42.1305 | 0 | 19.4414 | 18.0906 | 60.2495 | 24.8328 | 32.0936
| Rel.Std.Dev. |  | 1.427% | 2.078% | 2.609% |  | 2.411% | 1.464% | 1.204% | 1.174% | 3.318% | 0% | 1.293% | 1.208% | 4.113% | 1.691% | 2.161%
| PDM3 |  | -1.059% | -0.6694% | 3.391% |  | 1.611% | 1.725% | -0.3123% | 2.291% | -15.26% | -1.893% | 0.354% | -0.05024% | -2.227% | -2.004% | -0.8918%
| Z-Score (Absolute) |  | 0.24 | 0.08 | 1.65 | NA | 0.89 | 0.94 | 0.08 | 1.18 | 6.26 | 0.71 | 0.36 | 0.19 | 0.74 | 0.64 | 0.17
| Analytical results for Sn in OREAS 464 (Certified Value 45.9 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS |  | 1
2
| 1 | 1 | 48 | 45 | 48 | 45 | 46 | 46 | 36 | N.A. | 42 | 25.4136 | 50 | < 50 | 50 | 34 |  | 1
| 1 | 2 | 46 | 48 | 49 | 45 | 48 | 47 | 36 | N.A. | 43 | 22.4064 | 41 | < 50 | 50 | 33 |  | 10
| 1 | 3 | 49 | 46 | 47 | 46 | 46 | 46 | 37 | N.A. | 40 | 23.32 | 43 | < 50 | 50 | 33 |  | 16
| 1 | 4 | 49 | 44 | 48 | 45 | 46 | 46 | 37 | N.A. | 43 | 24.3963 | 31 | < 50 | 40 | 35 |  | 45.8667
| 1 | 5 | 47 | 45 | 48 | 46 | 44 | 46 | 36 | N.A. | 43 | 26.2892 | 43 | < 50 | 50 | 36 |  | 53
| 1 | 6 | 47 | 45 | 49 | 46 | 42 | 46 | 38 | N.A. | 43 | 25.8039 | 42 | < 50 | 50 | 35
| Mean |  | 47.6667 | 45.5 | 48.1667 | 45.5 | 45.3333 | 46.1667 | 36.6667 |  | 42.3333 | 24.6049 | 41.6667 |  | 48.3333 | 34.3333
| Median |  | 47.5 | 45 | 48 | 45.5 | 46 | 46 | 36.5 |  | 43 | 24.905 | 42.5 |  | 50 | 34.5
| Std Dev. |  | 1.2111 | 1.3784 | 0.75277 | 0.54772 | 2.0656 | 0.40825 | 0.8165 |  | 1.2111 | 1.5137 | 6.121 |  | 4.0825 | 1.2111
| Rel.Std.Dev. |  | 2.541% | 3.029% | 1.563% | 1.204% | 4.556% | 0.8843% | 2.227% |  | 2.861% | 6.152% | 14.69% |  | 8.447% | 3.527%
| PDM3 |  | 3.924% | -0.7994% | 5.015% | -0.7994% | -1.163% | 0.6541% | -20.06% |  | -7.703% | -46.36% | -9.157% |  | 5.378% | -25.15%
| Z-Score (Absolute) |  | 0.61 | 0.02 | 0.74 | 0.02 | 0.02 | 0.2 | 2.36 | NA | 0.83 | 5.61 | 0.44 | NA | 0.79 | 2.99
| Analytical results for Sr in OREAS 464 (Certified Value 2060 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 2150 | 1995 | 1985 | 2020 | 2030 | 2053 | 1997.5 | N.A. | 2100 | 2670.4789 | 1690 | 1930 | 2130 |  |  | 1
| 1 | 2 | 2110 | 2010 | 2080 | 2040 | 2090 | 2074 | 1985.6 | N.A. | 2100 | 2701.9895 | 1710 | 2030 | 2150 |  |  | 1
| 1 | 3 | 2240 | 2070 | 1985 | 2070 | 1995 | 2040 | 1970.4 | N.A. | 2000 | 2618.7746 | 1750 | 2040 | 2120 |  |  | 16
| 1 | 4 | 2220 | 2140 | 2070 | 2060 | 2020 | 2068 | 1986.2 | N.A. | 2100 | 2657.0456 | 1480 | 2020 | 2150 |  |  | 2060.41
| 1 | 5 | 2170 | 2080 | 2120 | 2060 | 2040 | 2075 | 1986.5 | N.A. | 2100 | 2714.4561 | 1700 | 1910 | 2160 |  |  | 54
| 1 | 6 | 2130 | 1910 | 2090 | 2030 | 1980 | 2065 | 2013.4 | N.A. | 2100 | 2741.1569 | 1660 | 2040 | 2140
| Mean |  | 2170 | 2034.1667 | 2055 | 2046.6667 | 2025.8333 | 2062.5 | 1989.9333 |  | 2083.3333 | 2683.9836 | 1665 | 1995 | 2141.6667
| Median |  | 2160 | 2040 | 2075 | 2050 | 2025 | 2066.5 | 1986.35 |  | 2100 | 2686.2342 | 1695 | 2025 | 2145
| Std Dev. |  | 50.9902 | 80.1509 | 56.745 | 19.6638 | 38.5249 | 13.5757 | 14.3805 |  | 40.8248 | 44.0042 | 95.2365 | 58.9067 | 14.7196
| Rel.Std.Dev. |  | 2.35% | 3.94% | 2.761% | 0.9608% | 1.902% | 0.6582% | 0.7227% |  | 1.96% | 1.64% | 5.72% | 2.953% | 0.6873%
| PDM3 |  | 5.319% | -1.274% | -0.2626% | -0.667% | -1.678% | 0.1014% | -3.421% |  | 1.113% | 30.26% | -19.19% | -3.175% | 3.944%
| Z-Score (Absolute) |  | 1.82 | 0.25 | 0.06 | 0.06 | 0.38 | 0.18 | 0.93 | NA | 0.5 | 9.67 | 5.33 | 0.85 | 1.39
| Analytical results for Ta in OREAS 464 (Certified Value 24.3 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 1
1
| 1 | 1 | 25 | 25.5 | 26.2 | 26.7 | 26 | 17.7509 | 19.8 | 485 | N.A. | 17 | 14.8768 | 19.5 | 19.5 | 25.3 |  | 1
| 1 | 2 | 24.4 | 25.2 | 26.8 | 27 | 25.8 | 18.0794 | 25.1 | 515 | N.A. | 19 | 10.5389 | 29.1 | 21 | 24.4 |  | 2
| 1 | 3 | 24.7 | 25.4 | 25.7 | 26.8 | 27 | 17.0517 | 25.5 | 430 | N.A. | 18 | 10.923 | 29.2 | 19.5 | 29.2 |  | 16
| 1 | 4 | 26.6 | 27.6 | 28.1 | 27.4 | 25.9 | 17.9976 | 25.5 | 533 | N.A. | 17 | 11.2654 | 21.9 | 22 | 24.9 |  | 24.3182
| 1 | 5 | 25.2 | 26 | 26.5 | 27.2 | 25.3 | 22.5212 | 25.6 | 524 | N.A. | 19 | 15.693 | 29.2 | 21.5 | 26 |  | 55
| 1 | 6 | 25 | 25.3 | 27.3 | 26.6 | 27 | 26.1362 | 25.4 | 541 | N.A. | 19 | 14.6487 | 28.6 | 22 | 25.3
| Mean |  | 25.15 | 25.8333 | 26.7667 | 26.95 | 26.1667 | 19.9228 | 24.4833 | 504.6667 |  | 18.1667 | 12.991 | 26.25 | 20.9167 | 25.85
| Median |  | 25 | 25.45 | 26.65 | 26.9 | 25.95 | 18.0385 | 25.45 | 519.5 |  | 18.5 | 12.9571 | 28.85 | 21.25 | 25.3
| Std Dev. |  | 0.76354 | 0.90921 | 0.84774 | 0.30822 | 0.68896 | 3.6172 | 2.3008 | 41.3892 |  | 0.98319 | 2.3183 | 4.3712 | 1.1583 | 1.7237
| Rel.Std.Dev. |  | 3.036% | 3.52% | 3.167% | 1.144% | 2.633% | 18.16% | 9.397% | 8.201% |  | 5.412% | 17.85% | 16.65% | 5.538% | 6.668%
| PDM3 |  | 3.42% | 6.23% | 10.07% | 10.82% | 7.601% | -18.07% | 0.6791% | 1975% |  | -25.3% | -46.58% | 7.944% | -13.99% | 6.299%
| Z-Score (Absolute) |  | 0.41 | 0 | 0.56 | 0.67 | 0.2 | 4.32 | 0.25 | 289.15 | NA | 4.63 | 7.75 | 1.07 | 2.97 | 0.01
| Analytical results for Tb in OREAS 464 (Certified Value 54 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 49.8 | 53 | 56.7 | 50 | 50.8 | 55.3537 | 52.4 | 53.2 | 45.0454 | 60 | 54 | 66.1319 | 48.8 | 54.6 | 51.2 | 1
| 1 | 2 | 49.7 | 52.1 | 55 | 50.8 | 52.4 | 55.0169 | 53.5 | 55.2 | 44.0845 | 60 | 55.1 | 66.5396 | 51.9 | 58.2 | 51.1 | 3
| 1 | 3 | 50.3 | 53 | 54.2 | 51.5 | 52.9 | 56.6655 | 52.8 | 52.6 | 45.4231 | 60 | 54.1 | 64.7892 | 53.4 | 55.6 | 47.2 | 16
| 1 | 4 | 52.2 | 54.4 | 56.5 | 50.8 | 52.6 | 55.5842 | 53.2 | 55.2 | 45.7544 | 60 | 54.9 | 64.7162 | 53 | 56.6 | 52 | 53.7883
| 1 | 5 | 51.1 | 52.3 | 55.9 | 50.7 | 50.5 | 58.1538 | 52.9 | 55.1 | 44.7563 | 60 | 55.7 | 65.9979 | 52.9 | 56.2 | 51.7 | 56
| 1 | 6 | 49.6 | 53.1 | 54.7 | 50.4 | 52.8 | 56.8095 | 53.4 | 53.9 | 45.2257 | 60 | 55.5 | 64.9279 | 50.6 | 56.2 | 52.6
| Mean |  | 50.45 | 52.9833 | 55.5 | 50.7 | 52 | 56.2639 | 53.0333 | 54.2 | 45.0482 | 60 | 54.8833 | 65.5171 | 51.7667 | 56.2333 | 50.9667
| Median |  | 50.05 | 53 | 55.45 | 50.75 | 52.5 | 56.1249 | 53.05 | 54.5 | 45.1356 | 60 | 55 | 65.4629 | 52.4 | 56.2 | 51.45
| Std Dev. |  | 1.0213 | 0.8085 | 1.0178 | 0.498 | 1.064 | 1.1727 | 0.41312 | 1.1367 | 0.58045 | 0 | 0.70545 | 0.79665 | 1.7694 | 1.1894 | 1.9253
| Rel.Std.Dev. |  | 2.024% | 1.526% | 1.834% | 0.9822% | 2.046% | 2.084% | 0.779% | 2.097% | 1.289% | 0% | 1.285% | 1.216% | 3.418% | 2.115% | 3.778%
| PDM3 |  | -6.206% | -1.497% | 3.182% | -5.742% | -3.325% | 4.603% | -1.404% | 0.7654% | -16.25% | 11.55% | 2.036% | 21.81% | -3.758% | 4.546% | -5.246%
| Z-Score (Absolute) |  | 0.75 | 0.01 | 0.71 | 0.67 | 0.3 | 0.93 | 0 | 0.34 | 2.31 | 1.87 | 0.53 | 3.61 | 0.37 | 0.92 | 0.38
| Analytical results for Te in OREAS 464 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 1 | < 1 | [13 empty] | 1
| 1 | 2 | < 1 | < 1 | [13 empty] | 16
| 1 | 3 | < 1 | < 1 | [13 empty] | 16
| 1 | 4 | < 1 | < 1 | [13 empty] | < 1
| 1 | 5 | < 1 | < 1 | [13 empty] | 22
| 1 | 6 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 464 (Certified Value 527 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 523 | 513 | 543 | 519 | 520 | 544.8934 | 550.4 | 527.6 | 366.6979 | 537 | 525.5 | 372.1322 | 417 | 503 | 493 | 1
| 1 | 2 | 519 | 521 | 540 | 531 | 535 | 541.0003 | 550.9 | 524.2 | 362.0363 | 550 | 520 | 377.414 | 473 | 585 | 481 | 0
| 1 | 3 | 538 | 532 | 520 | 539 | 504 | 550.6802 | 543.9 | 523.4 | 366.8701 | 549 | 515.5 | 372.2206 | 480 | 541 | 473 | 16
| 1 | 4 | 559 | 530 | 563 | 534 | 517 | 531.9659 | 539 | 536.1 | 366.8608 | 535 | 517.5 | 376.4704 | 498 | 567 | 493 | 527.4131
| 1 | 5 | 541 | 523 | 534 | 532 | 555 | 550.6077 | 550.5 | 525.8 | 368.495 | 543 | 526 | 381.6079 | 470 | 557 | 494 | 57
| 1 | 6 | 528 | 507 | 542 | 534 | 521 | 536.2771 | 556.1 | 521.6 | 367.4218 | 538 | 526 | 380.047 | 458 | 574 | 499
| Mean |  | 534.6667 | 521 | 540.3333 | 531.5 | 525.3333 | 542.5708 | 548.4667 | 526.45 | 366.397 | 542 | 521.75 | 376.6487 | 466 | 554.5 | 488.8333
| Median |  | 533 | 522 | 541 | 533 | 520.5 | 542.9468 | 550.45 | 525 | 366.8655 | 540.5 | 522.75 | 376.9422 | 471.5 | 562 | 493
| Std Dev. |  | 14.6242 | 9.654 | 13.9809 | 6.7157 | 17.5803 | 7.6196 | 6.0421 | 5.1528 | 2.2359 | 6.3875 | 4.6984 | 3.9183 | 27.3861 | 29.3513 | 9.7656
| Rel.Std.Dev. |  | 2.735% | 1.853% | 2.587% | 1.264% | 3.347% | 1.404% | 1.102% | 0.9788% | 0.6102% | 1.179% | 0.9005% | 1.04% | 5.877% | 5.293% | 1.998%
| PDM3 |  | 1.375% | -1.216% | 2.45% | 0.7749% | -0.3943% | 2.874% | 3.992% | -0.1826% | -30.53% | 2.766% | -1.074% | -28.59% | -11.64% | 5.136% | -7.315%
| Z-Score (Absolute) |  | 0.36 | 0.24 | 0.6 | 0.22 | 0.05 | 0.7 | 0.95 | 0 | 6.94 | 0.67 | 0.2 | 6.5 | 2.2 | 1.22 | 1.49
| Analytical results for Ti in OREAS 464 (Certified Value 1.95 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 1.983 | 1.94 | 1.85 | 1.9191 | N.A. | 2.01 | 1.8957 | 1.15 | 2.111 | 1.96 | 1.99 |  |  |  |  | 1
| 1 | 2 | 1.979 | 1.96 | 1.88 | 1.8432 | N.A. | 1.99 | 1.8997 | 1.83 | 2.056 | 1.92 | 1.97 |  |  |  |  | 4
| 1 | 3 | 2.005 | 1.95 | 1.89 | 1.8786 | N.A. | 1.98 | 1.8923 | 1.81 | 2.087 | 1.97 | 1.98 |  |  |  |  | 16
| 1 | 4 | 1.964 | 1.95 | 1.89 | 1.8457 | N.A. | 2 | 1.9176 | 1.31 | 2.029 | 1.96 | 1.98 |  |  |  |  | 1.9539
| 1 | 5 | 2.009 | 1.95 | 1.89 | 1.8647 | N.A. | 2 | 1.9049 | 1.84 | 2.001 | 2 | 1.97 |  |  |  |  | 58
| 1 | 6 | 2.002 | 1.93 | 1.89 | 1.8962 | N.A. | 2.02 | 1.902 | 1.85 | 2.014 | 1.98 | 1.96
| Mean |  | 1.9903 | 1.9467 | 1.8817 | 1.8746 |  | 2 | 1.902 | 1.6317 | 2.0497 | 1.965 | 1.975
| Median |  | 1.9925 | 1.95 | 1.89 | 1.8716 |  | 2 | 1.9009 | 1.82 | 2.0425 | 1.965 | 1.975
| Std Dev. |  | 0.01775 | 0.010328 | 0.016021 | 0.029579 |  | 0.014142 | 0.008841 | 0.31549 | 0.043043 | 0.026646 | 0.010488
| Rel.Std.Dev. |  | 0.8918% | 0.5305% | 0.8514% | 1.578% |  | 0.7071% | 0.4648% | 19.34% | 2.1% | 1.356% | 0.531%
| PDM3 |  | 1.865% | -0.3694% | -3.696% | -4.059% |  | 2.36% | -2.653% | -16.49% | 4.902% | 0.5689% | 1.081%
| Z-Score (Absolute) |  | 0.47 | 0.13 | 1.02 | 1.12 | NA | 0.61 | 0.74 | 3.14 | 1.29 | 0.13 | 0.26
| Analytical results for Tl in OREAS 464 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 14 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 17
| 1 | 3 | < 0.5 | N.A. | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | < 0.5
| 1 | 5 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 23
| 1 | 6 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5
| Mean |  |  |  | 0.1
| Median |  |  |  | 0.1
| Std Dev. |  |  |  | 0
| Rel.Std.Dev. |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Tm in OREAS 464 (Certified Value 3.56 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 3.4 | 3.5 | 3.57 | 3.8 | 3.45 | 3.3469 | 3.6 | 4.3 | 2.3656 | < 44 | 3.5 | 2.9158 | 3.47 | 3.4 | 3.33 | 1
| 1 | 2 | 3.38 | 3.59 | 3.71 | 3.82 | 3.39 | 3.2048 | 3.6 | 3.9 | 2.3018 | < 44 | 3.6 | 2.9578 | 3.73 | 3.6 | 3.57 | 5
| 1 | 3 | 3.39 | 3.49 | 3.72 | 3.81 | 3.65 | 3.5868 | 3.6 | 4 | 2.2367 | < 44 | 3.5 | 2.8878 | 3.81 | 3.6 | 3.28 | 16
| 1 | 4 | 3.51 | 3.56 | 3.9 | 3.86 | 3.38 | 3.3271 | 3.5 | 3.7 | 2.4043 | < 44 | 3.4 | 2.8954 | 3.67 | 3.6 | 3.42 | 3.5568
| 1 | 5 | 3.3 | 3.58 | 3.81 | 3.84 | 3.42 | 3.6116 | 3.6 | 4.2 | 2.3224 | < 44 | 3.5 | 2.9023 | 3.72 | 3.4 | 3.67 | 59
| 1 | 6 | 3.39 | 3.41 | 4.06 | 3.82 | 3.6 | 3.3298 | 3.6 | 3.8 | 2.3468 | < 44 | 3.4 | 2.8781 | 3.64 | 3.6 | 3.46
| Mean |  | 3.395 | 3.5217 | 3.795 | 3.825 | 3.4817 | 3.4012 | 3.5833 | 3.9833 | 2.3296 |  | 3.4833 | 2.9062 | 3.6733 | 3.5333 | 3.455
| Median |  | 3.39 | 3.53 | 3.765 | 3.82 | 3.435 | 3.3383 | 3.6 | 3.95 | 2.3346 |  | 3.5 | 2.8988 | 3.695 | 3.6 | 3.44
| Std Dev. |  | 0.067157 | 0.068532 | 0.17026 | 0.021679 | 0.11479 | 0.16176 | 0.040825 | 0.23166 | 0.057666 |  | 0.075277 | 0.028321 | 0.11535 | 0.10328 | 0.14625
| Rel.Std.Dev. |  | 1.978% | 1.946% | 4.487% | 0.5668% | 3.297% | 4.756% | 1.139% | 5.816% | 2.475% |  | 2.161% | 0.9745% | 3.14% | 2.923% | 4.233%
| PDM3 |  | -4.549% | -0.9879% | 6.697% | 7.54% | -2.112% | -4.376% | 0.7459% | 11.99% | -34.5% |  | -2.066% | -18.29% | 3.276% | -0.6599% | -2.862%
| Z-Score (Absolute) |  | 0.76 | 0.11 | 1.7 | 1.88 | 0.12 | 0.59 | 0.47 | 2.8 | 6.82 | NA | 0.11 | 3.47 | 0.99 | 0.18 | 0.28
| Analytical results for U in OREAS 464 (Certified Value 17.6 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 16.9 | 17.1 | 18 | 17.55 | 17.2 | 17.9064 | 17.5 | 17.8 | 14.1042 | 17 | 17.5 | 17.3384 | 9.93 | 15.5 | 17.9 | 1
| 1 | 2 | 17 | 17.3 | 17.7 | 17.45 | 17.3 | 18.2998 | 17.7 | 18.1 | 13.7472 | 17 | 17.5 | 17.5899 | 11.5 | 16 | 17.9 | 6
| 1 | 3 | 17 | 17.8 | 17.3 | 17.7 | 18.1 | 17.9553 | 17.2 | 17.8 | 14.5953 | 17 | 17.5 | 17.3219 | 11.7 | 15.5 | 16.7 | 16
| 1 | 4 | 17.2 | 17.9 | 18.2 | 17.7 | 17.3 | 17.6319 | 17.2 | 17.7 | 14.8109 | 17 | 18 | 17.5544 | 9.24 | 15.5 | 17.8 | 17.5665
| 1 | 5 | 16.6 | 17.3 | 17.5 | 18.1 | 16.8 | 18.4622 | 17.3 | 18.1 | 14.2086 | 17 | 18 | 17.6557 | 11.3 | 15.5 | 18.4 | 60
| 1 | 6 | 17.2 | 17.2 | 17.4 | 17.4 | 18.2 | 18.1336 | 17.7 | 17.7 | 14.7986 | 17 | 18 | 17.6021 | 11 | 15 | 17.7
| Mean |  | 16.9833 | 17.4333 | 17.6833 | 17.65 | 17.4833 | 18.0648 | 17.4333 | 17.8667 | 14.3775 | 17 | 17.75 | 17.5104 | 10.7783 | 15.5 | 17.7333
| Median |  | 17 | 17.3 | 17.6 | 17.625 | 17.3 | 18.0444 | 17.4 | 17.8 | 14.402 | 17 | 17.75 | 17.5722 | 11.15 | 15.5 | 17.85
| Std Dev. |  | 0.22286 | 0.33267 | 0.35449 | 0.25298 | 0.54924 | 0.29744 | 0.23381 | 0.18619 | 0.42736 | 0 | 0.27386 | 0.14345 | 0.97752 | 0.31623 | 0.56095
| Rel.Std.Dev. |  | 1.312% | 1.908% | 2.005% | 1.433% | 3.142% | 1.647% | 1.341% | 1.042% | 2.972% | 0% | 1.543% | 0.8192% | 9.069% | 2.04% | 3.163%
| PDM3 |  | -3.32% | -0.7583% | 0.6648% | 0.4751% | -0.4737% | 2.837% | -0.7583% | 1.708% | -18.15% | -3.225% | 1.044% | -0.3196% | -38.64% | -11.76% | 0.9494%
| Z-Score (Absolute) |  | 0.88 | 0.09 | 0.35 | 0.29 | 0 | 1.02 | 0.09 | 0.67 | 5.46 | 0.85 | 0.47 | 0.05 | 11.79 | 3.49 | 0.8
| Analytical results for V in OREAS 464 (Certified Value 231 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 222 | 238 | 247 | 232 | 210 | 236 | 224 | N.A. | 200 | 240.8823 | 263 | 250 | 215 |  |  | 1
| 1 | 2 | 215 | 236 | 242 | 231 | 218 | 240 | 225 | N.A. | 200 | 239.6509 | 263 | 200 | 210 |  |  | 7
| 1 | 3 | 228 | 236 | 245 | 237 | 210 | 246 | 220 | N.A. | 200 | 243.8919 | 264 | 250 | 204 |  |  | 16
| 1 | 4 | 230 | 259 | 249 | 237 | 214 | 238 | 216 | N.A. | 200 | 238.6089 | 275 | 250 | 214 |  |  | 230.9526
| 1 | 5 | 221 | 239 | 261 | 237 | 246 | 237 | 207 | N.A. | 200 | 242.1424 | 266 | 250 | 222 |  |  | 61
| 1 | 6 | 217 | 222 | 248 | 229 | 237 | 231 | 220 | N.A. | 200 | 247.0127 | 270 | 200 | 217
| Mean |  | 222.1667 | 238.3333 | 248.6667 | 233.8333 | 222.5 | 238 | 218.6667 |  | 200 | 242.0315 | 266.8333 | 233.3333 | 213.6667
| Median |  | 221.5 | 237 | 247.5 | 234.5 | 216 | 237.5 | 220 |  | 200 | 241.5123 | 265 | 250 | 214.5
| Std Dev. |  | 5.9133 | 11.8771 | 6.532 | 3.6009 | 15.2807 | 4.9396 | 6.5625 |  | 3.1134e-14 | 3.0655 | 4.7924 | 25.8199 | 6.1536
| Rel.Std.Dev. |  | 2.662% | 4.983% | 2.627% | 1.54% | 6.868% | 2.075% | 3.001% |  | 1.557e-14% | 1.267% | 1.796% | 11.07% | 2.88%
| PDM3 |  | -3.804% | 3.196% | 7.67% | 1.247% | -3.66% | 3.051% | -5.32% |  | -13.4% | 4.797% | 15.54% | 1.031% | -7.485%
| Z-Score (Absolute) |  | 0.68 | 0.04 | 0.76 | 0.01 | 0.66 | 0.26 | 0.89 | NA | 2.01 | 0.51 | 1.99 | 0.01 | 1.19
| Analytical results for W in OREAS 464 (Certified Value < 6 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 17 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 4 | 3 | 4 | 4 | 4 | 3 | 5 | < 10 | 5 | < 5 | 4 |  |  |  |  | 1
| 1 | 2 | 4 | 3 | 5 | 3 | 5 | 4 | 3 | < 10 | 6 | < 5 | 5 |  |  |  |  | 8
| 1 | 3 | 5 | 4 | 4 | 3 | 4 | 3 | 3 | < 10 | 6 | < 5 | 5 |  |  |  |  | 16
| 1 | 4 | 4 | 4 | 4 | 3 | 5 | 3 | 3 | 10 | 5 | < 5 | 5 |  |  |  |  | < 6
| 1 | 5 | 4 | 4 | 4 | 4 | 3 | 3 | 3 | 20 | 6 | < 5 | 5 |  |  |  |  | 62
| 1 | 6 | 3 | 4 | 5 | 3 | 3 | 3 | 3 | < 10 | 6 | < 5 | 5
| Mean |  | 4 | 3.6667 | 4.3333 | 3.3333 | 4 | 3.1667 | 3.3333 | 15 | 5.6667 |  | 4.8333
| Median |  | 4 | 4 | 4 | 3 | 4 | 3 | 3 | 15 | 6 |  | 5
| Std Dev. |  | 0.63246 | 0.5164 | 0.5164 | 0.5164 | 0.89443 | 0.40825 | 0.8165 | 7.0711 | 0.5164 |  | 0.40825
| Rel.Std.Dev. |  | 15.81% | 14.08% | 11.92% | 15.49% | 22.36% | 12.89% | 24.49% | 47.14% | 9.113% |  | 8.447%
PDM3
| Z-Score (Absolute) |  | 0 | 0.27 | 0.27 | 0.54 | 0 | 0.81 | 0.81 | 8.9 | 1.35 | NA | 0.81
| Analytical results for Y in OREAS 464 (Certified Value 449 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 441 | 453 | 461 | 458 | 448 | 493.1739 | 449 | 437.8 | 433 | 425 | 370.9 | 482 | 468 | 408 |  | 1
| 1 | 2 | 432 | 466 | 461 | 461 | 453 | 476.6405 | 455.3 | 433 | 425 | 432 | 381.1 | 572 | 494 | 402 |  | 0
| 1 | 3 | 444 | 463 | 457 | 463 | 464 | 467.9567 | 446.7 | 432.4 | 441 | 423.5 | 385.9 | 582 | 485 | 396 |  | 16
| 1 | 4 | 453 | 441 | 477 | 465 | 454 | 458.2024 | 452 | 431 | 425 | 419 | 378.8 | 474 | 473 | 406 |  | 448.5525
| 1 | 5 | 435 | 450 | 470 | 467 | 450 | 475.4697 | 454.9 | 470.7 | 433 | 435.5 | 389 | 580 | 468 | 416 |  | 63
| 1 | 6 | 439 | 447 | 466 | 459 | 476 | 473.925 | 451.9 | 442.1 | 433 | 438.5 | 385.1 | 568 | 482 | 415
| Mean |  | 440.6667 | 453.3333 | 465.3333 | 462.1667 | 457.5 | 474.228 | 451.6333 | 441.1667 | 431.6667 | 428.9167 | 381.8 | 543 | 478.3333 | 407.1667
| Median |  | 440 | 451.5 | 463.5 | 462 | 453.5 | 474.6974 | 451.95 | 435.4 | 433 | 428.5 | 383.1 | 570 | 477.5 | 407
| Std Dev. |  | 7.3937 | 9.5638 | 7.2847 | 3.4881 | 10.616 | 11.5113 | 3.3345 | 15.0422 | 6.0222 | 7.5856 | 6.4473 | 50.6715 | 10.4435 | 7.6529
| Rel.Std.Dev. |  | 1.678% | 2.11% | 1.565% | 0.7547% | 2.32% | 2.427% | 0.7383% | 3.41% | 1.395% | 1.769% | 1.689% | 9.332% | 2.183% | 1.88%
| PDM3 |  | -1.758% | 1.066% | 3.741% | 3.035% | 1.995% | 5.724% | 0.6868% | -1.647% | -3.765% | -4.378% | -14.88% | 21.06% | 6.639% | -9.227%
| Z-Score (Absolute) |  | 0.42 | 0.03 | 0.46 | 0.34 | 0.05 | 0.77 | 0.03 | 0.61 | 0.74 | 0.84 | 2.51 | 3.22 | 0.92 | 1.61
| Analytical results for Yb in OREAS 464 (Certified Value 15.7 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 15.7 | 15.15 | 14.7 | 15.5 | 16 | 14.1251 | 15.7 | 16.4 | 15.7738 | < 90 | 17.4 | 16.1761 | 15 | 14 | 19.9 | 1
| 1 | 2 | 15.6 | 15.7 | 14.4 | 15.3 | 16.3 | 15.8327 | 15.7 | 16.3 | 14.5855 | < 90 | 16.9 | 16.5412 | 17.9 | 15.5 | 19.2 | 9
| 1 | 3 | 15.6 | 16.05 | 14.6 | 15.7 | 15.4 | 15.5687 | 15.9 | 15.9 | 14.7507 | < 90 | 16.9 | 16.1267 | 18 | 16 | 19.1 | 16
| 1 | 4 | 16 | 17.5 | 14.8 | 15 | 15.9 | 14.7658 | 16.4 | 16.2 | 15.1834 | < 90 | 17.2 | 16.1324 | 15.2 | 15 | 19.7 | 15.7363
| 1 | 5 | 15 | 15.85 | 14.3 | 15.75 | 14.95 | 15.728 | 15.9 | 19.1 | 15.1046 | < 90 | 17.5 | 16.1791 | 18.5 | 15 | 20.1 | 64
| 1 | 6 | 15.2 | 15 | 15.4 | 15.6 | 14.05 | 14.0384 | 15.9 | 16.3 | 15.127 | < 90 | 17.7 | 15.9541 | 17.8 | 14 | 19.9
| Mean |  | 15.5167 | 15.875 | 14.7 | 15.475 | 15.4333 | 15.0098 | 15.9167 | 16.7 | 15.0875 |  | 17.2667 | 16.1849 | 17.0667 | 14.9167 | 19.65
| Median |  | 15.6 | 15.775 | 14.65 | 15.55 | 15.65 | 15.1672 | 15.9 | 16.3 | 15.1158 |  | 17.3 | 16.1543 | 17.85 | 15 | 19.8
| Std Dev. |  | 0.36009 | 0.89373 | 0.38987 | 0.2824 | 0.82926 | 0.81134 | 0.25626 | 1.1883 | 0.41169 |  | 0.3266 | 0.19314 | 1.5436 | 0.80104 | 0.40866
| Rel.Std.Dev. |  | 2.321% | 5.63% | 2.652% | 1.825% | 5.373% | 5.405% | 1.61% | 7.115% | 2.729% |  | 1.891% | 1.193% | 9.044% | 5.37% | 2.08%
| PDM3 |  | -1.396% | 0.8813% | -6.586% | -1.661% | -1.925% | -4.617% | 1.146% | 6.124% | -4.123% |  | 9.725% | 2.851% | 8.454% | -5.209% | 24.87%
| Z-Score (Absolute) |  | 0.18 | 0.25 | 1.15 | 0.23 | 0.28 | 0.78 | 0.18 | 0.66 | 0.69 | NA | 1.9 | 0.62 | 1.67 | 0.9 | 4.74
| Analytical results for Zn in OREAS 464 (Certified Value 1232 ppm) | [16 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 14 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | N.A. | 1200 | 1351.5745 | 1280 | 1150 | 1130 | [9 empty] | 1
| 1 | 2 | N.A. | 1200 | 1365.8146 | 1250 | 1150 | 1120 | [9 empty] | 10
| 1 | 3 | N.A. | 1200 | 1401.9443 | 1270 | 1200 | 1110 | [9 empty] | 16
| 1 | 4 | N.A. | 1200 | 1424.4096 | 1280 | 1200 | 1110 | [9 empty] | 1231.6406
| 1 | 5 | N.A. | 1200 | 1389.3696 | 1230 | 1250 | 1110 | [9 empty] | 65
| 1 | 6 | N.A. | 1200 | 1406.1048 | 1210 | 1250 | 1110
| Mean |  |  | 1200 | 1389.8696 | 1253.3333 | 1200 | 1115
| Median |  |  | 1200 | 1395.657 | 1260 | 1200 | 1110
| Std Dev. |  |  | 2.4908e-13 | 27.0104 | 28.7518 | 44.7214 | 8.3666
| Rel.Std.Dev. |  |  | 2.076e-14% | 1.943% | 2.294% | 3.727% | 0.7504%
| PDM3 |  |  | -2.569% | 12.85% | 1.761% | -2.569% | -9.47%
| Z-Score (Absolute) |  | NA | 0 | 2.4 | 0.67 | 0 | 1.07
| Analytical results for Zr in OREAS 464 (Certified Value 462 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 460 | 460 | 470 | 466 | 460 | 496 | 442 | N.A. | 446 | 445.7 | 1100 |  |  |  |  | 1
| 1 | 2 | 440 | 470 | 480 | 462 | 450 | 486 | 441 | N.A. | 445 | 355.3 | 1960 |  |  |  |  | 11
| 1 | 3 | 470 | 460 | 487 | 467 | 450 | 475 | 425 | N.A. | 436 | 330.2 | 1970 |  |  |  |  | 16
| 1 | 4 | 480 | 460 | 486 | 478 | 450 | 477 | 434 | N.A. | 445 | 346.5 | 1080 |  |  |  |  | 461.7292
| 1 | 5 | 450 | 460 | 493 | 470 | 450 | 477 | 439 | N.A. | 443 | 370.2 | 1990 |  |  |  |  | 66
| 1 | 6 | 460 | 470 | 498 | 467 | 470 | 466 | 447 | N.A. | 449 | 300.6 | 1900
| Mean |  | 460 | 463.3333 | 485.6667 | 468.3333 | 455 | 479.5 | 438 |  | 444 | 358.0833 | 1666.6667
| Median |  | 460 | 460 | 486.5 | 467 | 450 | 477 | 440 |  | 445 | 350.9 | 1930
| Std Dev. |  | 14.1421 | 5.164 | 9.8522 | 5.3914 | 8.3666 | 10.2908 | 7.642 |  | 4.3818 | 49.0861 | 447.735
| Rel.Std.Dev. |  | 3.074% | 1.115% | 2.029% | 1.151% | 1.839% | 2.146% | 1.745% |  | 0.9869% | 13.71% | 26.86%
| PDM3 |  | -0.3745% | 0.3474% | 5.184% | 1.43% | -1.457% | 3.849% | -5.139% |  | -3.84% | -22.45% | 261%
| Z-Score (Absolute) |  | 0.06 | 0.06 | 0.91 | 0.25 | 0.25 | 0.68 | 0.9 | NA | 0.67 | 3.94 | 45.78
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 464 (Certified Value 4.43 ppm) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 5.11 | 5.07 | 2.82 | 3.98 | 3.64 | 3.9122 | 4.3 | 4.8 | < 0.01 | 4.95 | 4 | 4.5 |  |  | 1
| 1 | 2 | 4.5 | 4.91 | 1.72 | 3.84 | 3.2 | 4.2153 | 4.3 | 4.5 | < 0.01 | 4.62 | 4 | 4.5 |  |  | 15
| 1 | 3 | 4.25 | 4.27 | 2.55 | 3.96 | 3.26 | 4.1407 | 4.7 | 4.5 | < 0.01 | 5.03 | 4 | 4.5 |  |  | 16
| 1 | 4 | 4.67 | 4.43 | 1.94 | 4.09 | 3.3 | 4.2482 | 4.4 | 4.9 | < 0.01 | 4.88 | 6 | 4.5 |  |  | 4.431
| 1 | 5 | 4.66 | 4.67 | 3.03 | 4.19 | 3.35 | 3.9543 | 4.2 | 4.5 | < 0.01 | 4.96 | 4 | 4.5 |  |  | 68
| 1 | 6 | 4.84 | 4.85 | 3.39 | 3.91 | 3.28 | 4.1732 | 4.1 | 5 | < 0.01 | 4.79 | 4 | 4.5
| Mean |  | 4.6717 | 4.7 | 2.575 | 3.995 | 3.3383 | 4.1073 | 4.3333 | 4.7 |  | 4.8717 | 4.3333 | 4.5
| Median |  | 4.665 | 4.76 | 2.685 | 3.97 | 3.29 | 4.1569 | 4.3 | 4.65 |  | 4.915 | 4 | 4.5
| Std Dev. |  | 0.29253 | 0.30404 | 0.64283 | 0.12629 | 0.15575 | 0.14031 | 0.20656 | 0.22804 |  | 0.1477 | 0.8165 | 0
| Rel.Std.Dev. |  | 6.262% | 6.469% | 24.96% | 3.161% | 4.665% | 3.416% | 4.767% | 4.852% |  | 3.032% | 18.84% | 0%
| PDM3 |  | 5.431% | 6.071% | -41.89% | -9.84% | -24.66% | -7.305% | -2.204% | 6.071% |  | 9.945% | -2.204% | 1.557%
| Z-Score (Absolute) |  | 0.67 | 0.73 | 3.5 | 0.67 | 2.1 | 0.45 | 0 | 0.73 | NA | 1.07 | 0.66 | 0.33
| Analytical results for Al in OREAS 464 (Certified Value 4.47 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 4.31 | 4.47 | 4.55 | 4.82 | 4.38 | 4.4999 | 4.45 | 4.58 | 4.45 | 4.5104 | 4.1042 | 4.61 | 4.576 | 4.7 | 1
| 1 | 2 | 4.33 | 4.42 | 4.49 | 4.81 | 4.61 | 4.4001 | 4.51 | 4.55 | 4.39 | 4.6046 | 4.1117 | 4.79 | 4.583 | 4.65 | 1
| 1 | 3 | 4.25 | 4.42 | 4.49 | 4.77 | 4.27 | 4.4034 | 4.54 | 4.55 | 4.48 | 4.4891 | 4.1659 | 4.65 | 4.528 | 4.83 | 16
| 1 | 4 | 4.28 | 4.37 | 4.49 | 4.97 | 4.61 | 4.5123 | 4.47 | 4.58 | 4.38 | 4.3918 | 4.1773 | 4.89 | 4.553 | 4.78 | 4.47
| 1 | 5 | 4.31 | 4.47 | 4.53 | 4.91 | 4.5 | 4.4466 | 4.51 | 4.47 | 4.42 | 4.3676 | 4.1645 | 4.7 | 4.586 | 4.7 | 69
| 1 | 6 | 4.37 | 4.49 | 4.56 | 4.79 | 4.38 | 4.6427 | 4.48 | 4.61 | 4.45 | 4.3029 | 4.1643 | 4.82 | 4.582 | 4.81
| Mean |  | 4.3083 | 4.44 | 4.5183 | 4.845 | 4.4583 | 4.4841 | 4.4933 | 4.5567 | 4.4283 | 4.4444 | 4.148 | 4.7433 | 4.568 | 4.745
| Median |  | 4.31 | 4.445 | 4.51 | 4.815 | 4.44 | 4.4732 | 4.495 | 4.565 | 4.435 | 4.4405 | 4.1644 | 4.745 | 4.579 | 4.74
| Std Dev. |  | 0.041191 | 0.044721 | 0.032506 | 0.07791 | 0.13819 | 0.090754 | 0.03266 | 0.048028 | 0.038687 | 0.11019 | 0.031476 | 0.10764 | 0.022952 | 0.071764
| Rel.Std.Dev. |  | 0.9561% | 1.007% | 0.7194% | 1.608% | 3.1% | 2.024% | 0.7269% | 1.054% | 0.8736% | 2.479% | 0.7588% | 2.269% | 0.5025% | 1.512%
| PDM3 |  | -3.616% | -0.6709% | 1.082% | 8.39% | -0.2607% | 0.3168% | 0.5223% | 1.939% | -0.9319% | -0.5724% | -7.204% | 6.115% | 2.193% | 6.152%
| Z-Score (Absolute) |  | 1.9 | 0.51 | 0.31 | 3.74 | 0.32 | 0.05 | 0.05 | 0.71 | 0.63 | 0.47 | 3.58 | 2.68 | 0.83 | 2.69
| Analytical results for As in OREAS 464 (Certified Value 32.9 ppm) | [15 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 36.2 | 35.8 | 28.1 | 34.4 | 52.6 | 32.0107 | 28 | 37 | 32.1 | 84.5 | < 50 | 70 |  |  | 1
| 1 | 2 | 35.6 | 35 | 29.1 | 32.3 | 51.8 | 30.2213 | 29 | 39 | 30.9 | 77.6 | < 50 | 74 |  |  | 16
| 1 | 3 | 36.3 | 31.6 | 27.5 | 32.4 | 52.5 | 32.2249 | 27 | 46 | 31.9 | 80.2 | < 50 | 72 |  |  | 16
| 1 | 4 | 36 | 28.5 | 28.9 | 31.9 | 54.9 | 30.5103 | 29 | 43 | 32 | 80.1 | < 50 | 69 |  |  | 32.9362
| 1 | 5 | 35.4 | 34 | 27 | 34.1 | 54.8 | 29.6059 | 28 | 44 | 30.9 | 82.7 | < 50 | 73 |  |  | 70
| 1 | 6 | 36.7 | 32.5 | 26.8 | 33.8 | 51.4 | 32.2638 | 29 | 46 | 31.4 | 80 | < 50 | 72
| Mean |  | 36.0333 | 32.9 | 27.9 | 33.15 | 53 | 31.1395 | 28.3333 | 42.5 | 31.5333 | 80.85 |  | 71.6667
| Median |  | 36.1 | 33.25 | 27.8 | 33.1 | 52.55 | 31.2605 | 28.5 | 43.5 | 31.65 | 80.15 |  | 72
| Std Dev. |  | 0.4761 | 2.6548 | 0.9654 | 1.071 | 1.5007 | 1.1655 | 0.8165 | 3.7283 | 0.5465 | 2.4089 |  | 1.8619
| Rel.Std.Dev. |  | 1.321% | 8.069% | 3.46% | 3.231% | 2.831% | 3.743% | 2.882% | 8.772% | 1.733% | 2.98% |  | 2.598%
| PDM3 |  | 9.403% | -0.1099% | -15.29% | 0.6492% | 60.92% | -5.455% | -13.98% | 29.04% | -4.259% | 145.5% |  | 117.6%
| Z-Score (Absolute) |  | 0.4 | 0.03 | 0.73 | 0 | 2.78 | 0.28 | 0.67 | 1.31 | 0.23 | 6.68 | NA | 5.39
| Analytical results for Ba in OREAS 464 (Certified Value 1648 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 1650 | 1730 | 1700 | 1770 | 1630 | 1672.0617 | 1509 | 1663 | 1650 | 1580.4 | 1636.93 | 1495 | 1655 | 1570 | 1
| 1 | 2 | 1640 | 1680 | 1690 | 1780 | 1720 | 1730.6561 | 1525 | 1662 | 1600 | 1529.1 | 1623.25 | 1583 | 1665 | 1590 | 2
| 1 | 3 | 1650 | 1700 | 1680 | 1760 | 1590 | 1708.8388 | 1511 | 1667 | 1650 | 1700.8 | 1641.51 | 1553 | 1655 | 1590 | 16
| 1 | 4 | 1660 | 1680 | 1720 | 1840 | 1700 | 1696.2428 | 1525 | 1646 | 1600 | 1565.4 | 1653.42 | 1649 | 1655 | 1610 | 1648.1311
| 1 | 5 | 1640 | 1400 | 1710 | 1820 | 1660 | 1697.2616 | 1497 | 1638 | 1670 | 2797.8 | 1634.77 | 1566 | 1660 | 1640 | 71
| 1 | 6 | 1670 | 1730 | 1700 | 1780 | 1640 | 1665.6526 | 1484 | 1670 | 1680 | 1558.8 | 1651.02 | 1618 | 1655 | 1630
| Mean |  | 1651.6667 | 1653.3333 | 1700 | 1791.6667 | 1656.6667 | 1695.1189 | 1508.5 | 1657.6667 | 1641.6667 | 1788.7167 | 1640.15 | 1577.3333 | 1657.5 | 1605
| Median |  | 1650 | 1690 | 1700 | 1780 | 1650 | 1696.7522 | 1510 | 1662.5 | 1650 | 1572.9 | 1639.22 | 1574.5 | 1655 | 1600
| Std Dev. |  | 11.6905 | 126.1216 | 14.1421 | 31.2517 | 47.6095 | 23.9035 | 16.0219 | 12.7227 | 34.3026 | 497.897 | 11.1447 | 53.5188 | 4.1833 | 26.6458
| Rel.Std.Dev. |  | 0.7078% | 7.628% | 0.8319% | 1.744% | 2.874% | 1.41% | 1.062% | 0.7675% | 2.089% | 27.84% | 0.6795% | 3.393% | 0.2524% | 1.66%
| PDM3 |  | 0.2145% | 0.3156% | 3.147% | 8.709% | 0.5179% | 2.851% | -8.472% | 0.5786% | -0.3922% | 8.53% | -0.4843% | -4.296% | 0.5685% | -2.617%
| Z-Score (Absolute) |  | 0.04 | 0.77 | 0.71 | 2.14 | 0.04 | 0.64 | 2.26 | 0.05 | 0.19 | 1.05 | 0.22 | 1.19 | 0.05 | 0.76
| Analytical results for Be in OREAS 464 (Certified Value 15.2 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 13.05 | 15.35 | 13.6 | 17.15 | 11.45 | 15.3338 | 15.2 | 16.5 | 15.1 | 15.06 | 13.5936 | 19 | 15.6 |  | 1
| 1 | 2 | 13.25 | 15.6 | 13.6 | 16.6 | 10.75 | 15.38 | 15.7 | 16.5 | 15.2 | 15.47 | 14.1298 | 19 | 16 |  | 0
| 1 | 3 | 13.4 | 13.85 | 14.7 | 17.15 | 10.85 | 15.3764 | 15.2 | 17.6 | 15.4 | 15.84 | 14.1889 | 19 | 16 |  | 16
| 1 | 4 | 13.8 | 14.3 | 13.9 | 17.8 | 10.6 | 15.12 | 15.3 | 17.3 | 14.9 | 15.91 | 13.9878 | 19.5 | 15.7 |  | 15.2434
| 1 | 5 | 13.6 | 14.45 | 14.05 | 17.25 | 10.75 | 15.0475 | 14.9 | 17.1 | 15.3 | 16.68 | 14.0684 | 19 | 16.4 |  | 72
| 1 | 6 | 13.95 | 15.35 | 13.6 | 16.95 | 10.2 | 15.1839 | 14.9 | 17 | 15.5 | 15.32 | 14.0353 | 19.5 | 16.2
| Mean |  | 13.5083 | 14.8167 | 13.9083 | 17.15 | 10.7667 | 15.2403 | 15.2 | 17 | 15.2333 | 15.7133 | 14.0006 | 19.1667 | 15.9833
| Median |  | 13.5 | 14.9 | 13.75 | 17.15 | 10.75 | 15.2588 | 15.2 | 17.05 | 15.25 | 15.655 | 14.0518 | 19 | 16
| Std Dev. |  | 0.33973 | 0.70969 | 0.43176 | 0.3937 | 0.40579 | 0.14254 | 0.29665 | 0.43818 | 0.21602 | 0.57095 | 0.21156 | 0.2582 | 0.29944
| Rel.Std.Dev. |  | 2.515% | 4.79% | 3.104% | 2.296% | 3.769% | 0.9353% | 1.952% | 2.578% | 1.418% | 3.634% | 1.511% | 1.347% | 1.873%
| PDM3 |  | -11.38% | -2.799% | -8.758% | 12.51% | -29.37% | -0.02055% | -0.2847% | 11.52% | -0.06599% | 3.083% | -8.153% | 25.74% | 4.854%
| Z-Score (Absolute) |  | 1.01 | 0.24 | 0.87 | 1.12 | 2.62 | 0 | 0.02 | 1.03 | 0 | 0.28 | 0.67 | 2.3 | 0.44
| Analytical results for Bi in OREAS 464 (Certified Value 4.05 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.29 | 4.07 | 4.17 | 4.18 | 3.63 | 3.897 | 4.2 | 4.4 | < 0.01 | 4.12 | 3.8264 | 3.6 | 3.9 |  | 1
| 1 | 2 | 4.14 | 4.06 | 4.11 | 4.1 | 3.58 | 3.9294 | 4.3 | 4.4 | < 0.01 | 4.2 | 3.5306 | 3.6 | 3.9 |  | 0
| 1 | 3 | 4.26 | 3.65 | 4.1 | 4.16 | 3.6 | 3.8733 | 4.2 | 4.5 | < 0.01 | 4.26 | 3.6509 | 3.6 | 4 |  | 16
| 1 | 4 | 4.27 | 4.01 | 4.05 | 4.13 | 3.66 | 3.899 | 4.3 | 4.4 | < 0.01 | 4.1 | 3.6172 | 3.5 | 4 |  | 4.0538
| 1 | 5 | 4.23 | 4.05 | 4.27 | 4.3 | 3.74 | 3.7816 | 4.2 | 4.3 | < 0.01 | 4.24 | 3.6808 | 3.5 | 4.1 |  | 73
| 1 | 6 | 4.31 | 4.09 | 4.06 | 4.19 | 3.66 | 3.8252 | 4.1 | 4.7 | < 0.01 | 4.15 | 3.8453 | 3.5 | 4
| Mean |  | 4.25 | 3.9883 | 4.1267 | 4.1767 | 3.645 | 3.8676 | 4.2167 | 4.45 |  | 4.1783 | 3.6919 | 3.55 | 3.9833
| Median |  | 4.265 | 4.055 | 4.105 | 4.17 | 3.645 | 3.8852 | 4.2 | 4.4 |  | 4.175 | 3.6659 | 3.55 | 4
| Std Dev. |  | 0.060332 | 0.16786 | 0.082138 | 0.068896 | 0.05648 | 0.054592 | 0.075277 | 0.13784 |  | 0.065243 | 0.12248 | 0.054772 | 0.075277
| Rel.Std.Dev. |  | 1.42% | 4.209% | 1.99% | 1.65% | 1.55% | 1.412% | 1.785% | 3.098% |  | 1.561% | 3.318% | 1.543% | 1.89%
| PDM3 |  | 4.839% | -1.616% | 1.797% | 3.03% | -10.09% | -4.594% | 4.017% | 9.773% |  | 3.071% | -8.929% | -12.43% | -1.739%
| Z-Score (Absolute) |  | 0.75 | 0.17 | 0.17 | 0.41 | 2.12 | 1.06 | 0.6 | 1.47 | NA | 0.41 | 1.9 | 2.57 | 0.51
| Analytical results for Ca in OREAS 464 (Certified Value 0.851 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.84 | 0.82 | 0.38 | 0.42 | 0.36 | 0.79497 | 0.82 | 0.92 | 0.75 | 0.9246 | 0.80297 | 0.84 | 0.877 | 0.95 | 1
| 1 | 2 | 0.84 | 0.83 | 0.37 | 0.41 | 0.36 | 0.79579 | 0.83 | 0.92 | 0.746 | 0.9413 | 0.79702 | 0.89 | 0.877 | 0.93 | 3
| 1 | 3 | 0.83 | 0.83 | 0.38 | 0.42 | 0.37 | 0.78263 | 0.83 | 0.91 | 0.745 | 0.9202 | 0.80992 | 0.87 | 0.8695 | 0.97 | 16
| 1 | 4 | 0.85 | 0.85 | 0.38 | 0.43 | 0.38 | 0.81991 | 0.82 | 0.92 | 0.748 | 0.8969 | 0.81383 | 0.87 | 0.8765 | 0.96 | 0.85095
| 1 | 5 | 0.83 | 0.81 | 0.37 | 0.43 | 0.36 | 0.77293 | 0.83 | 0.91 | 0.744 | 0.8886 | 0.80815 | 0.86 | 0.878 | 0.93 | 74
| 1 | 6 | 0.84 | 0.89 | 0.38 | 0.43 | 0.37 | 0.84059 | 0.82 | 0.93 | 0.734 | 0.8795 | 0.81121 | 0.87 | 0.8785 | 0.94
| Mean |  | 0.83833 | 0.83833 | 0.37667 | 0.42333 | 0.36667 | 0.80114 | 0.825 | 0.91833 | 0.7445 | 0.90852 | 0.80718 | 0.86667 | 0.87608 | 0.94667
| Median |  | 0.84 | 0.83 | 0.38 | 0.425 | 0.365 | 0.79538 | 0.825 | 0.92 | 0.7455 | 0.90855 | 0.80904 | 0.87 | 0.877 | 0.945
| Std Dev. |  | 0.0075277 | 0.028577 | 0.005164 | 0.008165 | 0.008165 | 0.024955 | 0.0054772 | 0.0075277 | 0.0055767 | 0.023847 | 0.0061628 | 0.01633 | 0.0033078 | 0.01633
| Rel.Std.Dev. |  | 0.8979% | 3.409% | 1.371% | 1.929% | 2.227% | 3.115% | 0.6639% | 0.8197% | 0.7491% | 2.625% | 0.7635% | 1.884% | 0.3776% | 1.725%
| PDM3 |  | -1.482% | -1.482% | -55.74% | -50.25% | -56.91% | -5.853% | -3.049% | 7.919% | -12.51% | 6.765% | -5.143% | 1.847% | 2.954% | 11.25%
| Z-Score (Absolute) |  | 0.12 | 0.02 | 4.61 | 4.13 | 4.71 | 0.26 | 0.02 | 0.94 | 0.84 | 0.84 | 0.2 | 0.41 | 0.51 | 1.23
| Analytical results for Cd in OREAS 464 (Certified Value 0.52 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.55 | 0.54 | 0.46 | 0.45 | 0.43 | 0.53871 | 0.7 | 0.86 | 0.51 | 0.5 | 1.5 |  |  |  | 1
| 1 | 2 | 0.53 | 0.6 | 0.55 | 0.48 | 0.46 | 0.55078 | 0.6 | 0.84 | 0.38 | 0.5 | 1 |  |  |  | 18
| 1 | 3 | 0.58 | 0.59 | 0.51 | 0.52 | 0.42 | 0.55 | 0.7 | 0.88 | 0.35 | 0.5 | 1 |  |  |  | 16
| 1 | 4 | 0.55 | 0.63 | 0.44 | 0.54 | 0.43 | 0.59346 | 0.7 | 0.8 | 0.29 | 0.5 | 1 |  |  |  | 0.51813
| 1 | 5 | 0.5 | 0.58 | 0.58 | 0.49 | 0.43 | 0.59238 | 0.5 | 0.86 | 0.42 | 0.5 | 1 |  |  |  | 75
| 1 | 6 | 0.5 | 0.68 | 0.54 | 0.52 | 0.44 | 0.56365 | 0.6 | 0.82 | 0.35 | < 0.5 | 1
| Mean |  | 0.535 | 0.60333 | 0.51333 | 0.5 | 0.435 | 0.56483 | 0.63333 | 0.84333 | 0.38333 | 0.5 | 1.0833
| Median |  | 0.54 | 0.595 | 0.525 | 0.505 | 0.43 | 0.55722 | 0.65 | 0.85 | 0.365 | 0.5 | 1
| Std Dev. |  | 0.031464 | 0.04761 | 0.054283 | 0.032863 | 0.013784 | 0.023149 | 0.08165 | 0.029439 | 0.075277 | 0 | 0.20412
| Rel.Std.Dev. |  | 5.881% | 7.891% | 10.57% | 6.573% | 3.169% | 4.098% | 12.89% | 3.491% | 19.64% | 0% | 18.84%
| PDM3 |  | 3.256% | 16.44% | -0.9257% | -3.499% | -16.04% | 9.014% | 22.23% | 62.76% | -26.02% | -3.499% | 109.1%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.21 | 0.35 | 1.04 | 0.29 | 0.97 | 3.04 | 1.5 | 0.35 | 4.59
| Analytical results for Ce in OREAS 464 (Certified Value 1.55 wt.%) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5828 | 0.83941 | 1.5021 | > 0.5 | 1.4678 | 1.6123 | 1.5599 | 1.52 |  | 1
| 1 | 2 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5715 | 0.85922 | 1.5313 | > 0.5 | 1.4714 | 1.5574 | 1.5737 | 1.54 |  | 0
| 1 | 3 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5691 | 0.89569 | 1.5263 | > 0.5 | 1.4926 | 1.5379 | 1.6387 | 1.55 |  | 16
| 1 | 4 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5636 | 0.83316 | 1.4983 | > 0.5 | 1.4801 | 1.5788 | 1.5558 | 1.56 |  | 1.5483
| 1 | 5 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5773 | 0.76918 | 1.4883 | > 0.5 | 1.5131 | 1.5943 | 1.5812 | 1.57 |  | 76
| 1 | 6 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 1.5554 | 0.87641 | 1.5078 | > 0.5 | 1.4616 | 1.6522 | 1.6453 | 1.55
| Mean |  | 07: 1.5699; 08: 0.84551; 09: 1.509; 11: 1.4811; 12: 1.5888; 17: 1.5924; 22: 1.5483
| Median |  | 07: 1.5703; 08: 0.84932; 09: 1.5049; 11: 1.4758; 12: 1.5866; 17: 1.5775; 22: 1.55
| Std Dev. |  | 07: 0.0097359; 08: 0.043993; 09: 0.016682; 11: 0.019021; 12: 0.040684; 17: 0.03951; 22: 0.017224
| Rel.Std.Dev. |  | 07: 0.6201%; 08: 5.203%; 09: 1.106%; 11: 1.284%; 12: 2.561%; 17: 2.481%; 22: 1.112%
| PDM3 |  | 07: 1.399%; 08: -45.39%; 09: -2.537%; 11: -4.338%; 12: 2.619%; 17: 2.853%; 22: 0.003829%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.36 | 11.71 | 0.66 | NA | 1.12 | 0.67 | 0.73 | 0
| Analytical results for Co in OREAS 464 (Certified Value 17 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 16.7 | 16.6 | 16.4 | 17.4 | 16.7 | 16.4107 | 17.9 | 19 | 16.7 | 17.5 | 17.1681 | 25 | 15 |  | 1
| 1 | 2 | 16.3 | 16.8 | 16.4 | 17 | 16.6 | 16.0219 | 17.8 | 19 | 16.5 | 16.8 | 17.6635 | 25 | 15 |  | 0
| 1 | 3 | 16.7 | 14.7 | 16.9 | 17.1 | 16.2 | 16.236 | 17.6 | 18 | 16.9 | 17.7 | 17.245 | 25 | 15 |  | 16
| 1 | 4 | 17 | 15.8 | 16.6 | 17.2 | 16.6 | 16.0575 | 17.7 | 18 | 16.4 | 17 | 17.2167 | 25 | 10 |  | 16.9903
| 1 | 5 | 16.5 | 16.8 | 16.4 | 17.6 | 16.9 | 16.261 | 17.1 | 17 | 16.9 | 17.7 | 17.3671 | 25 | 20 |  | 77
| 1 | 6 | 16.9 | 16.4 | 16 | 17.2 | 16.4 | 16.5666 | 17.4 | 19 | 16.9 | 17.1 | 16.905 | 25 | 20
| Mean |  | 16.6833 | 16.1833 | 16.45 | 17.25 | 16.5667 | 16.259 | 17.5833 | 18.3333 | 16.7167 | 17.3 | 17.2609 | 25 | 15.8333
| Median |  | 16.7 | 16.5 | 16.4 | 17.2 | 16.6 | 16.2485 | 17.65 | 18.5 | 16.8 | 17.3 | 17.2309 | 25 | 15
| Std Dev. |  | 0.25626 | 0.81588 | 0.29496 | 0.21679 | 0.24221 | 0.20734 | 0.29269 | 0.8165 | 0.22286 | 0.38471 | 0.24934 | 0 | 3.7639
| Rel.Std.Dev. |  | 1.536% | 5.042% | 1.793% | 1.257% | 1.462% | 1.275% | 1.665% | 4.454% | 1.333% | 2.224% | 1.445% | 0% | 23.77%
| PDM3 |  | -1.807% | -4.749% | -3.18% | 1.529% | -2.493% | -4.304% | 3.491% | 7.905% | -1.61% | 1.823% | 1.593% | 47.14% | -6.809%
| Z-Score (Absolute) |  | 0.49 | 0.81 | 0.86 | 0.39 | 0.67 | 1.15 | 0.91 | 2.07 | 0.44 | 0.47 | 0.41 | 12.45 | 0
| Analytical results for Cr in OREAS 464 (Certified Value 335 ppm) | [15 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 309 | 310 | 143 | 159 | 140 | 322.3758 | 337 | 337 | 390 | 278 | 400 |  |  |  | 1
| 1 | 2 | 310 | 306 | 139 | 153 | 141 | 330.0602 | 343 | 329 | 394 | 285 | 400 |  |  |  | 4
| 1 | 3 | 304 | 307 | 143 | 158 | 143 | 330.818 | 353 | 320 | 378 | 276 | 400 |  |  |  | 16
| 1 | 4 | 314 | 314 | 139 | 159 | 145 | 322.9994 | 344 | 319 | 372 | 294 | 400 |  |  |  | 334.6808
| 1 | 5 | 303 | 279 | 139 | 164 | 139 | 326.0025 | 347 | 318 | 366 | 282 | 400 |  |  |  | 78
| 1 | 6 | 311 | 332 | 144 | 164 | 148 | 335.4223 | 357 | 334 | 358 | 287 | 400
| Mean |  | 308.5 | 308 | 141.1667 | 159.5 | 142.6667 | 327.9464 | 346.8333 | 326.1667 | 376.3333 | 283.6667 | 400
| Median |  | 309.5 | 308.5 | 141 | 159 | 142 | 328.0314 | 345.5 | 324.5 | 375 | 283.5 | 400
| Std Dev. |  | 4.2308 | 17.0997 | 2.4014 | 4.1352 | 3.3862 | 5.0574 | 7.2226 | 8.2805 | 13.8804 | 6.532 | 0
| Rel.Std.Dev. |  | 1.371% | 5.552% | 1.701% | 2.593% | 2.374% | 1.542% | 2.082% | 2.539% | 3.688% | 2.303% | 0%
| PDM3 |  | -7.823% | -7.972% | -57.82% | -52.34% | -57.37% | -2.012% | 3.631% | -2.544% | 12.45% | -15.24% | 19.52%
| Z-Score (Absolute) |  | 0 | 0.01 | 2.94 | 2.62 | 2.92 | 0.34 | 0.67 | 0.31 | 1.19 | 0.44 | 1.61
| Analytical results for Cs in OREAS 464 (Certified Value 0.29 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.28 | 0.28 | 0.3 | 0.28 | 0.26 | 0.25138 | 0.3 | < 0.5 | 0.3 | 0.13 | 0.36023 | < 0.5 | 0.3 |  | 1
| 1 | 2 | 0.29 | 0.3 | 0.28 | 0.27 | 0.26 | 0.2379 | 0.3 | < 0.5 | 0.31 | 0.14 | 0.39273 | < 0.5 | 0.2 |  | 0
| 1 | 3 | 0.27 | 0.3 | 0.3 | 0.28 | 0.26 | 0.27407 | 0.3 | < 0.5 | 0.32 | 0.14 | 0.41295 | < 0.5 | 0.3 |  | 16
| 1 | 4 | 0.29 | 0.28 | 0.3 | 0.29 | 0.27 | 0.25584 | 0.3 | < 0.5 | 0.3 | 0.15 | 0.35685 | < 0.5 | 0.3 |  | 0.28775
| 1 | 5 | 0.29 | 0.51 | 0.31 | 0.3 | 0.27 | 0.27302 | 0.3 | < 0.5 | 0.31 | 0.17 | 0.38254 | < 0.5 | 0.3 |  | 79
| 1 | 6 | 0.29 | 0.27 | 0.3 | 0.3 | 0.27 | 0.27025 | 0.3 | < 0.5 | 0.31 | 0.15 | 0.33117 | < 0.5 | 0.3
| Mean |  | 0.285 | 0.32333 | 0.29833 | 0.28667 | 0.265 | 0.26041 | 0.3 |  | 0.30833 | 0.14667 | 0.37275 |  | 0.28333
| Median |  | 0.29 | 0.29 | 0.3 | 0.285 | 0.265 | 0.26304 | 0.3 |  | 0.31 | 0.145 | 0.37139 |  | 0.3
| Std Dev. |  | 0.0083666 | 0.092232 | 0.0098319 | 0.012111 | 0.0054772 | 0.014499 | 0 |  | 0.0075277 | 0.013663 | 0.029152 |  | 0.040825
| Rel.Std.Dev. |  | 2.936% | 28.53% | 3.296% | 4.225% | 2.067% | 5.568% | 0% |  | 2.441% | 9.315% | 7.821% |  | 14.41%
| PDM3 |  | -0.9554% | 12.37% | 3.678% | -0.3762% | -7.906% | -9.501% | 4.257% |  | 7.154% | -49.03% | 29.54% |  | -1.535%
| Z-Score (Absolute) |  | 0.08 | 0.03 | 0.59 | 0 | 1.1 | 1.33 | 0.67 | NA | 1.1 | 7.08 | 4.35 | NA | 0.67
| Analytical results for Cu in OREAS 464 (Certified Value 92 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 92.9 | 94.2 | 89.5 | 92.3 | 97.3 | 95.0027 | 92.9 | 108 | 89.3 | 148 | 92.172 | 81.4 | 90 | 122 | 1
| 1 | 2 | 92.2 | 94.4 | 89.3 | 90.4 | 93.2 | 93.9913 | 94.1 | 111 | 89.1 | 142 | 91.9861 | 90.4 | 90 | 120 | 19
| 1 | 3 | 92.7 | 88.9 | 89.9 | 90.9 | 91.4 | 93.7953 | 93 | 117 | 89.1 | 144 | 92.2955 | 87.6 | 90 | 126 | 16
| 1 | 4 | 96 | 93.3 | 89.8 | 91 | 93.8 | 95.0528 | 93.5 | 97 | 91.7 | 138 | 92.3413 | 88.9 | 90 | 126 | 91.8599
| 1 | 5 | 92.1 | 93.7 | 90.2 | 93.4 | 94.7 | 95.2364 | 91.1 | 88 | 90 | 135 | 92.3954 | 87.4 | 90 | 124 | 80
| 1 | 6 | 93.4 | 96.5 | 87.7 | 90.8 | 93.5 | 95.1866 | 90.9 | 122 | 90.3 | 134 | 92.1176 | 91.2 | 90 | 124
| Mean |  | 93.2167 | 93.5 | 89.4 | 91.4667 | 93.9833 | 94.7109 | 92.5833 | 107.1667 | 89.9167 | 140.1667 | 92.218 | 87.8167 | 90 | 123.6667
| Median |  | 92.8 | 93.95 | 89.65 | 90.95 | 93.65 | 95.0277 | 92.95 | 109.5 | 89.65 | 140 | 92.2338 | 88.25 | 90 | 124
| Std Dev. |  | 1.4442 | 2.5116 | 0.88994 | 1.1448 | 1.9529 | 0.64196 | 1.2999 | 12.6715 | 1.0048 | 5.4559 | 0.15405 | 3.4839 | 0 | 2.3381
| Rel.Std.Dev. |  | 1.549% | 2.686% | 0.9955% | 1.252% | 2.078% | 0.6778% | 1.404% | 11.82% | 1.118% | 3.892% | 0.167% | 3.967% | 0% | 1.891%
| PDM3 |  | 1.477% | 1.785% | -2.678% | -0.4281% | 2.312% | 3.104% | 0.7875% | 16.66% | -2.115% | 52.59% | 0.3898% | -4.402% | -2.025% | 34.63%
| Z-Score (Absolute) |  | 0.01 | 0.51 | 0.92 | 0.33 | 0.39 | 0.6 | 0.01 | 4.16 | 0.77 | 13.61 | 0.12 | 1.01 | 0.75 | 8.89
| Analytical results for Dy in OREAS 464 (Certified Value 182 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 177 | 177.5 | 172.5 | 180.5 | 180.5 | 189.7844 | 139.8 | 184.9 | 105 | 179.04 | 186.7164 | 178.5 | 191 |  | 1
| 1 | 2 | 177.5 | 175.5 | 172 | 178 | 177.5 | 184.9931 | 148 | 187.4 | 102 | 181.34 | 186.618 | 178.8 | 192 |  | 5
| 1 | 3 | 177 | 162 | 173 | 179 | 178.5 | 186.4419 | 147.1 | 189.8 | 111 | 187.28 | 186.5737 | 184 | 198 |  | 16
| 1 | 4 | 174.5 | 171 | 172.5 | 180 | 182.5 | 187.0597 | 144.1 | 185.7 | 105 | 179.32 | 187.3401 | 176.9 | 191 |  | 181.7073
| 1 | 5 | 177 | 170.5 | 173.5 | 186.5 | 189 | 188.9754 | 138.4 | 183.8 | 107 | 188.38 | 185.8073 | 180.2 | 197 |  | 81
| 1 | 6 | 176 | 177 | 169.5 | 180.5 | 183.5 | 183.0957 | 144.8 | 188.6 | 113 | 179.99 | 186.4277 | 184.3 | 191
| Mean |  | 176.5 | 172.25 | 172.1667 | 180.75 | 181.9167 | 186.725 | 143.7 | 186.7 | 107.1667 | 182.5583 | 186.5805 | 180.45 | 193.3333
| Median |  | 177 | 173.25 | 172.5 | 180.25 | 181.5 | 186.7508 | 144.45 | 186.55 | 106 | 180.665 | 186.5959 | 179.5 | 191.5
| Std Dev. |  | 1.0954 | 5.8374 | 1.4024 | 2.9791 | 4.1523 | 2.4814 | 3.8657 | 2.2961 | 4.1191 | 4.1745 | 0.49356 | 3.0534 | 3.266
| Rel.Std.Dev. |  | 0.6206% | 3.389% | 0.8145% | 1.648% | 2.283% | 1.329% | 2.69% | 1.23% | 3.844% | 2.287% | 0.2645% | 1.692% | 1.689%
| PDM3 |  | -2.866% | -5.205% | -5.251% | -0.5269% | 0.1152% | 2.761% | -20.92% | 2.748% | -41.02% | 0.4683% | 2.682% | -0.692% | 6.398%
| Z-Score (Absolute) |  | 0.43 | 0.88 | 0.89 | 0.09 | 0.16 | 0.68 | 3.96 | 0.67 | 7.91 | 0.23 | 0.66 | 0 | 1.39
| Analytical results for Er in OREAS 464 (Certified Value 36.9 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 39.4 | 37.7 | 43.9 | 36.9 | 39.2 | 37.1888 | 29.4 | 36.9 | 34.7 | 34.66 | 68.2043 | 36.3 | 36.5 |  | 1
| 1 | 2 | 38.5 | 37.4 | 43.3 | 36.3 | 38.2 | 37.4241 | 31.6 | 36.9 | 33.9 | 35.12 | 66.8857 | 36.75 | 37 |  | 6
| 1 | 3 | 38.9 | 33.3 | 43.2 | 36.5 | 39.4 | 36.2596 | 31.2 | 36.6 | 35.5 | 35 | 66.0024 | 38.05 | 39.2 |  | 16
| 1 | 4 | 38.3 | 36 | 42.5 | 36.2 | 39.3 | 36.7 | 30.3 | 37.3 | 33.6 | 34.84 | 66.889 | 36.35 | 37.7 |  | 36.9362
| 1 | 5 | 38.7 | 36.5 | 44.2 | 37.9 | 40.5 | 37.2366 | 29.8 | 36.2 | 34.6 | 36.21 | 65.9028 | 36.8 | 38.2 |  | 82
| 1 | 6 | 38.8 | 37.8 | 43.4 | 36.5 | 39.5 | 36.2438 | 30.4 | 37.5 | 34.5 | 34.72 | 66.8975 | 37.6 | 37.2
| Mean |  | 38.7667 | 36.45 | 43.4167 | 36.7167 | 39.35 | 36.8422 | 30.45 | 36.9 | 34.4667 | 35.0917 | 66.7969 | 36.975 | 37.6333
| Median |  | 38.75 | 36.95 | 43.35 | 36.5 | 39.35 | 36.9444 | 30.35 | 36.9 | 34.55 | 34.92 | 66.8874 | 36.775 | 37.45
| Std Dev. |  | 0.37771 | 1.6979 | 0.59133 | 0.62743 | 0.73417 | 0.51615 | 0.82885 | 0.46904 | 0.66533 | 0.57405 | 0.82921 | 0.70338 | 0.96471
| Rel.Std.Dev. |  | 0.9743% | 4.658% | 1.362% | 1.709% | 1.866% | 1.401% | 2.722% | 1.271% | 1.93% | 1.636% | 1.241% | 1.902% | 2.563%
| PDM3 |  | 4.956% | -1.316% | 17.55% | -0.5943% | 6.535% | -0.2546% | -17.56% | -0.09796% | -6.686% | -4.994% | 80.84% | 0.1051% | 1.887%
| Z-Score (Absolute) |  | 0.67 | 0.04 | 2.42 | 0.19 | 0.89 | 0.05 | 2.46 | 0.03 | 0.94 | 0.79 | 11.22 | 0 | 0.25
| Analytical results for Eu in OREAS 464 (Certified Value 334 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 376 | 327 | 330 | 344 | 342 | 338.8633 | 249.2 | 345.4 | 158 | 317.68 | 336.1098 | 318.8 | 302 |  | 1
| 1 | 2 | 377 | 327 | 326 | 341 | 338 | 333.4651 | 267.9 | 352.5 | 155 | 312.15 | 335.3974 | 319.9 | 300 |  | 7
| 1 | 3 | 379 | 297 | 331 | 343 | 341 | 325.6455 | 261 | 352.8 | 150 | 323.6 | 335.5798 | 331 | 311 |  | 16
| 1 | 4 | 377 | 314 | 325 | 341 | 349 | 322.9648 | 250 | 347.5 | 154 | 314.69 | 336.8207 | 315.4 | 309 |  | 334.4206
| 1 | 5 | 374 | 321 | 324 | 354 | 356 | 336.0578 | 238.7 | 348.9 | 153 | 325.38 | 337.7347 | 320.7 | 311 |  | 83
| 1 | 6 | 376 | 330 | 321 | 345 | 344 | 327.1213 | 255.1 | 353.3 | 159 | 311.25 | 339.152 | 330.3 | 307
| Mean |  | 376.5 | 319.3333 | 326.1667 | 344.6667 | 345 | 330.6863 | 253.65 | 350.0667 | 154.8333 | 317.4583 | 336.7991 | 322.6833 | 306.6667
| Median |  | 376.5 | 324 | 325.5 | 343.5 | 343 | 330.2932 | 252.55 | 350.7 | 154.5 | 316.185 | 336.4652 | 320.3 | 308
| Std Dev. |  | 1.6432 | 12.3396 | 3.7639 | 4.8442 | 6.5115 | 6.3432 | 10.1569 | 3.2733 | 3.3116 | 5.9145 | 1.438 | 6.4341 | 4.6762
| Rel.Std.Dev. |  | 0.4364% | 3.864% | 1.154% | 1.405% | 1.887% | 1.918% | 4.004% | 0.9351% | 2.139% | 1.863% | 0.427% | 1.994% | 1.525%
| PDM3 |  | 12.58% | -4.511% | -2.468% | 3.064% | 3.163% | -1.117% | -24.15% | 4.679% | -53.7% | -5.072% | 0.7112% | -3.51% | -8.299%
| Z-Score (Absolute) |  | 2.04 | 0.1 | 0 | 0.67 | 0.76 | 0.18 | 2.94 | 0.97 | 6.95 | 0.35 | 0.43 | 0.14 | 0.79
| Analytical results for Fe in OREAS 464 (Certified Value 35.55 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 33.6 | 35.7 | 34.5 | 37.6 | 33.8 | 38.257 | 34.13 | 37.32 | 35.4 | 36.4 | 33.192 | > 15 | 36.1085 | 36.3 | 1
| 1 | 2 | 33.7 | 35.3 | 34.3 | 37.6 | 35.4 | 38.3827 | 34.41 | 35.97 | 35.4 | 37.27 | 32.9459 | > 15 | 36.261 | 35.9 | 0
| 1 | 3 | 33.2 | 35.2 | 33.9 | 37.5 | 33.2 | 38.5951 | 34.65 | 36.12 | 35.4 | 36.22 | 33.2037 | > 15 | 36.048 | 37.3 | 16
| 1 | 4 | 33.6 | 34.6 | 34.7 | 38.9 | 35.4 | 38.3439 | 34.16 | 36.66 | 35.9 | 35.1 | 33.4527 | > 15 | 36.2385 | 36.8 | 35.551
| 1 | 5 | 33.3 | 34.9 | 34.6 | 38.9 | 34.7 | 37.6751 | 34.36 | 35.51 | 35.3 | 35.03 | 33.4198 | > 15 | 36.185 | 36.8 | 84
| 1 | 6 | 34.4 | 35.9 | 34.5 | 37.5 | 34.1 | 37.8 | 34.29 | 37.28 | 35.6 | 34.68 | 33.4838 | > 15 | 36.2235 | 37
| Mean |  | 33.6333 | 35.2667 | 34.4167 | 38 | 34.4333 | 38.1756 | 34.3333 | 36.4767 | 35.5 | 35.7833 | 33.283 |  | 36.1774 | 36.6833
| Median |  | 33.6 | 35.25 | 34.5 | 37.6 | 34.4 | 38.3005 | 34.325 | 36.39 | 35.4 | 35.66 | 33.3118 |  | 36.2043 | 36.8
| Std Dev. |  | 0.42269 | 0.48442 | 0.28577 | 0.69857 | 0.89144 | 0.35929 | 0.18981 | 0.73601 | 0.21909 | 1.0033 | 0.20786 |  | 0.082942 | 0.50365
| Rel.Std.Dev. |  | 1.257% | 1.374% | 0.8303% | 1.838% | 2.589% | 0.9411% | 0.5528% | 2.018% | 0.6172% | 2.804% | 0.6245% |  | 0.2293% | 1.373%
| PDM3 |  | -5.394% | -0.7997% | -3.191% | 6.889% | -3.144% | 7.383% | -3.425% | 2.604% | -0.1434% | 0.6536% | -6.38% |  | 1.762% | 3.185%
| Z-Score (Absolute) |  | 1.16 | 0.15 | 0.67 | 1.56 | 0.66 | 1.67 | 0.73 | 0.61 | 0 | 0.18 | 1.38 | NA | 0.42 | 0.74
| Analytical results for Ga in OREAS 464 (Indicative Value 97 ppm) | [15 empty] | Indicative
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 10; 08: 12; 09: 22 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS | 0
0
| 1 | 1 | 01: 109.5; 02: 83; 03: 93.4; 05: 160.5; 06: 37.5221; 07: < 0.05; 08: 103.9513; 09: FL | 1
| 1 | 2 | 01: 112.5; 02: 82.6; 03: 92.2; 05: 159.5; 06: 35.6445; 07: < 0.05; 08: 103.9887; 09: FL | 21
| 1 | 3 | 01: 102; 02: 83.6; 03: 92.5; 05: 155; 06: 39.7219; 07: < 0.05; 08: 103.2342; 09: FL | 16
| 1 | 4 | 01: 105; 02: 84.2; 03: 91.7; 05: 160; 06: 38.5406; 07: < 0.05; 08: 106.35; 09: FL | 97.0371
| 1 | 5 | 01: 108; 02: 81.7; 03: 96; 05: 161; 06: 36.2705; 07: < 0.05; 08: 106.8086; 09: FL | 27
| 1 | 6 | 01: 113; 02: 80.1; 03: 91.9; 05: 156; 06: 40.2133; 07: < 0.05; 08: 105.3179; 09: FL
| Mean |  | 01: 108.3333; 02: 82.5333; 03: 92.95; 05: 158.6667; 06: 37.9855; 08: 104.9418
| Median |  | 01: 108.75; 02: 82.8; 03: 92.35; 05: 159.75; 06: 38.0313; 08: 104.6533
| Std Dev. |  | 01: 4.2856; 02: 1.4665; 03: 1.6084; 05: 2.5232; 06: 1.8397; 08: 1.4432
| Rel.Std.Dev. |  | 01: 3.956%; 02: 1.777%; 03: 1.73%; 05: 1.59%; 06: 4.843%; 08: 1.375%
| PDM3 |  | 01: 11.64%; 02: -14.95%; 03: -4.212%; 05: 63.51%; 06: -60.85%; 08: 8.146%
| Z-Score (Absolute) |  | 01: 0.51; 02: 0.84; 03: 0.33; 05: 3.14; 06: 3.17; 07: NA; 08: 0.33; 09: NA
| Analytical results for Gd in OREAS 464 (Certified Value 670 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 690 | 679 | 661 | 610 | 689 | 683.1972 | 535.3 | 690.9 | 142 | 667.82 | 843.5711 | 660.5 | 643 |  | 1
| 1 | 2 | 693 | 668 | 656 | 603 | 680 | 675.2026 | 532.3 | 708.3 | 146 | 670.04 | 843.0187 | 659 | 640 |  | 8
| 1 | 3 | 687 | 595 | 658 | 617 | 680 | 657.3935 | 548.1 | 706 | 144 | 671.69 | 841.5674 | 686.3 | 660 |  | 16
| 1 | 4 | 693 | 670 | 655 | 602 | 693 | 669.2507 | 545.2 | 706.7 | 147 | 659.72 | 842.0982 | 655.6 | 669 |  | 670.037
| 1 | 5 | 699 | 703 | 663 | 636 | 715 | 690.2606 | 519.3 | 694.3 | 151 | 677.27 | 844.6831 | 670.3 | 678 |  | 85
| 1 | 6 | 685 | 692 | 646 | 608 | 693 | 661.6857 | 550.2 | 708.7 | 146 | 658.29 | 842.0767 | 690.4 | 655
| Mean |  | 691.1667 | 667.8333 | 656.5 | 612.6667 | 691.6667 | 672.8317 | 538.4 | 702.4833 | 146 | 667.4717 | 842.8359 | 670.35 | 657.5
| Median |  | 691.5 | 674.5 | 657 | 609 | 691 | 672.2266 | 540.25 | 706.35 | 146 | 668.93 | 842.5584 | 665.4 | 657.5
| Std Dev. |  | 4.9967 | 38.0916 | 5.9582 | 12.6438 | 12.8634 | 12.5875 | 11.7576 | 7.7945 | 3.0332 | 7.2789 | 1.1589 | 14.8311 | 14.7071
| Rel.Std.Dev. |  | 0.7229% | 5.704% | 0.9076% | 2.064% | 1.86% | 1.871% | 2.184% | 1.11% | 2.078% | 1.091% | 0.1375% | 2.212% | 2.237%
| PDM3 |  | 3.154% | -0.3289% | -2.02% | -8.562% | 3.228% | 0.4171% | -19.65% | 4.842% | -78.21% | -0.3829% | 25.79% | 0.04671% | -1.871%
| Z-Score (Absolute) |  | 0.67 | 0.39 | 0.45 | 2.02 | 0.69 | 0.08 | 4.27 | 1.04 | 16.99 | 0.09 | 5.59 | 0 | 0.42
| Analytical results for Ge in OREAS 464 (Indicative Value 18.4 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 11.65 | 18.45 | 16.5 | 15 | 33.7 | 0.5 | 14.6 | 34 | < 0.05 | 21.9 |  |  |  |  | 1
| 1 | 2 | 11.3 | 17.95 | 16.75 | 15 | 33.1 | 0.5 | 15.7 | 34 | < 0.05 | 21 |  |  |  |  | 15
| 1 | 3 | 11.7 | 16.05 | 16.7 | 14.9 | 31.9 | 0.4 | 15.5 | 36.8 | < 0.05 | 21.57 |  |  |  |  | 16
| 1 | 4 | 11.8 | 16.9 | 16.15 | 14.8 | 33 | 0.4 | 14.2 | 36.1 | < 0.05 | 22.41 |  |  |  |  | 18.3781
| 1 | 5 | 11.45 | 17.25 | 16.25 | 15.45 | 34 | 0.5693 | 13.5 | 35.1 | < 0.05 | 21.24 |  |  |  |  | 27
| 1 | 6 | 11.45 | 17.6 | 16 | 15.05 | 32.8 | 0.50585 | 15.2 | 35.1 | < 0.05 | 21.02
| Mean |  | 11.5583 | 17.3667 | 16.3917 | 15.0333 | 33.0833 | 0.47919 | 14.7833 | 35.1833 |  | 21.5233
| Median |  | 11.55 | 17.425 | 16.375 | 15 | 33.05 | 0.5 | 14.9 | 35.1 |  | 21.405
| Std Dev. |  | 0.18819 | 0.84063 | 0.30564 | 0.22286 | 0.73598 | 0.066691 | 0.84242 | 1.1197 |  | 0.55457
| Rel.Std.Dev. |  | 1.628% | 4.841% | 1.865% | 1.482% | 2.225% | 13.92% | 5.698% | 3.182% |  | 2.577%
| PDM3 |  | -37.11% | -5.503% | -10.81% | -18.2% | 80.02% | -97.39% | -19.56% | 91.44% |  | 17.11%
| Z-Score (Absolute) |  | 0.67 | 0.14 | 0 | 0.19 | 2.33 | 2.22 | 0.22 | 2.62 | NA | 0.72
| Analytical results for Hf in OREAS 464 (Certified Value 7.79 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 2
2
| 1 | 1 | 9.5 | 6.9 | 8.8 | 7.7 | 7.9 | 5.6756 | 6.9 | 1.18 | 7.81 | 10 |  |  |  |  | 1
| 1 | 2 | 9.2 | 7 | 8.9 | 7.8 | 7.6 | 6.0314 | 6.3 | 1.01 | 7.28 | 10 |  |  |  |  | 9
| 1 | 3 | 9.2 | 6.2 | 8.6 | 7.8 | 7.7 | 5.5216 | 6.2 | 1.08 | 7.91 | 10 |  |  |  |  | 16
| 1 | 4 | 9.3 | 6.6 | 8.7 | 7.9 | 7.8 | 5.7156 | 6.4 | 1.08 | 7.95 | 10 |  |  |  |  | 7.7917
| 1 | 5 | 8.9 | 6.1 | 9 | 8 | 8 | 5.4057 | 6.4 | 1.16 | 7.6 | 10 |  |  |  |  | 86
| 1 | 6 | 9.2 | 7 | 8.6 | 7.6 | 7.8 | 6.1802 | 6.8 | 1.14 | 7.37 | 10
| Mean |  | 9.2167 | 6.6333 | 8.7667 | 7.8 | 7.8 | 5.755 | 6.5 | 1.1083 | 7.6533 | 10
| Median |  | 9.2 | 6.75 | 8.75 | 7.8 | 7.8 | 5.6956 | 6.4 | 1.11 | 7.705 | 10
| Std Dev. |  | 0.19408 | 0.40332 | 0.1633 | 0.14142 | 0.14142 | 0.29717 | 0.28284 | 0.063377 | 0.28317 | 0
| Rel.Std.Dev. |  | 2.106% | 6.08% | 1.863% | 1.813% | 1.813% | 5.164% | 4.351% | 5.718% | 3.7% | 0%
| PDM3 |  | 18.29% | -14.87% | 12.51% | 0.1069% | 0.1069% | -26.14% | -16.58% | -85.78% | -1.775% | 28.34%
| Z-Score (Absolute) |  | 0.87 | 0.64 | 0.6 | 0.04 | 0.04 | 1.15 | 0.71 | 3.85 | 0.04 | 1.32
| Analytical results for Ho in OREAS 464 (Certified Value 21.3 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 23.2 | 20.1 | 21.8 | 20.8 | 21.7 | 20.9285 | 17 | 20.8 | 7.41 | 20.28 | 26.7814 | 21 | 21.4 |  | 1
| 1 | 2 | 23.3 | 21.4 | 21.7 | 20.7 | 21.3 | 20.5754 | 17.9 | 21.5 | 7.86 | 19.85 | 25.8772 | 21.15 | 22.4 |  | 10
| 1 | 3 | 23.3 | 19.8 | 21.5 | 20.8 | 21.6 | 20.5605 | 17.6 | 21.1 | 7.6 | 20.57 | 25.1493 | 21.8 | 21.8 |  | 16
| 1 | 4 | 23.1 | 21 | 21.6 | 20.6 | 21.9 | 20.4383 | 17.6 | 21.2 | 7.28 | 20.27 | 25.5899 | 20.6 | 21.8 |  | 21.3127
| 1 | 5 | 23.1 | 19.65 | 21.9 | 21.7 | 22.7 | 20.82 | 16.7 | 20.9 | 7.58 | 20.9 | 25.9525 | 21.1 | 22.9 |  | 87
| 1 | 6 | 23.4 | 21 | 21.3 | 20.9 | 22 | 20.4192 | 17.4 | 21.1 | 7.42 | 19.84 | 26.9036 | 21.55 | 21.3
| Mean |  | 23.2333 | 20.4917 | 21.6333 | 20.9167 | 21.8667 | 20.6237 | 17.3667 | 21.1 | 7.525 | 20.285 | 26.0423 | 21.2 | 21.9333
| Median |  | 23.25 | 20.55 | 21.65 | 20.8 | 21.8 | 20.5679 | 17.5 | 21.1 | 7.5 | 20.275 | 25.9149 | 21.125 | 21.8
| Std Dev. |  | 0.12111 | 0.7324 | 0.21602 | 0.39707 | 0.4761 | 0.20687 | 0.44121 | 0.24495 | 0.20256 | 0.41157 | 0.68199 | 0.42308 | 0.6121
| Rel.Std.Dev. |  | 0.5213% | 3.574% | 0.9986% | 1.898% | 2.177% | 1.003% | 2.541% | 1.161% | 2.692% | 2.029% | 2.619% | 1.996% | 2.791%
| PDM3 |  | 9.012% | -3.852% | 1.504% | -1.858% | 2.599% | -3.233% | -18.51% | -0.998% | -64.69% | -4.822% | 22.19% | -0.5288% | 2.912%
| Z-Score (Absolute) |  | 1.88 | 0.54 | 0.47 | 0.3 | 0.67 | 0.42 | 3.28 | 0 | 11.94 | 0.72 | 4.35 | 0.09 | 0.73
| Analytical results for In in OREAS 464 (Certified Value 2.11 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 2.17 | 2.15 | 2.46 | 2.15 | 2.17 | 2.0199 | 2.3 | 2.2 | 2.12 | 2 | 2.1 | 1.88 |  |  | 1
| 1 | 2 | 2.16 | 2.09 | 2.44 | 2.12 | 2.17 | 1.991 | 2.3 | 2.28 | 2.13 | 2.01 | 2.1 | 1.8 |  |  | 0
| 1 | 3 | 2.21 | 1.905 | 2.42 | 2.09 | 2.12 | 1.9906 | 2.3 | 2.22 | 2.14 | 2.01 | 2.2 | 1.94 |  |  | 16
| 1 | 4 | 2.2 | 1.99 | 2.43 | 2.18 | 2.2 | 2.0049 | 2.3 | 2.13 | 2.16 | 2.01 | 2.1 | 1.84 |  |  | 2.1115
| 1 | 5 | 2.18 | 2.05 | 2.55 | 2.32 | 2.28 | 1.9995 | 2.3 | 2.19 | 2.16 | 2.17 | 2.2 | 1.96 |  |  | 88
| 1 | 6 | 2.23 | 2.09 | 2.47 | 2.12 | 2.18 | 2.015 | 2.2 | 2.23 | 2.15 | 2.03 | 2.1 | 1.9
| Mean |  | 2.1917 | 2.0458 | 2.4617 | 2.1633 | 2.1867 | 2.0035 | 2.2833 | 2.2083 | 2.1433 | 2.0383 | 2.1333 | 1.8867
| Median |  | 2.19 | 2.07 | 2.45 | 2.135 | 2.175 | 2.0022 | 2.3 | 2.21 | 2.145 | 2.01 | 2.1 | 1.89
| Std Dev. |  | 0.026394 | 0.086857 | 0.047081 | 0.082624 | 0.052789 | 0.012177 | 0.040825 | 0.049565 | 0.01633 | 0.065243 | 0.05164 | 0.060222
| Rel.Std.Dev. |  | 1.204% | 4.246% | 1.913% | 3.819% | 2.414% | 0.6078% | 1.788% | 2.244% | 0.7619% | 3.201% | 2.421% | 3.192%
| PDM3 |  | 3.796% | -3.11% | 16.58% | 2.454% | 3.559% | -5.117% | 8.137% | 4.585% | 1.507% | -3.466% | 1.033% | -10.65%
| Z-Score (Absolute) |  | 0.44 | 0.77 | 2.68 | 0.05 | 0.4 | 1.12 | 1.2 | 0.58 | 0.04 | 1.05 | 0.04 | 2.09
| Analytical results for K in OREAS 464 (Certified Value 0.081 wt.%) | [15 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.07 | 0.09 | 0.09 | 0.07 | 0.1 | 0.1 | 0.09 | 0.089 | 0.0695 | 0.07 | 0.07 | 0.06 |  |  | 1
| 1 | 2 | 0.07 | 0.09 | 0.09 | 0.07 | 0.1 | 0.1 | 0.08 | 0.088 | 0.0737 | 0.08 | 0.07 | 0.06 |  |  | 11
| 1 | 3 | 0.07 | 0.1 | 0.09 | 0.07 | 0.09 | 0.1 | 0.09 | 0.089 | 0.0711 | 0.08 | 0.07 | 0.06 |  |  | 16
| 1 | 4 | 0.07 | 0.11 | 0.09 | 0.07 | 0.1 | 0.1 | 0.08 | 0.086 | 0.0644 | 0.08 | 0.07 | 0.06 |  |  | 0.081232
| 1 | 5 | 0.07 | 0.19 | 0.09 | 0.07 | 0.1 | 0.1 | 0.07 | 0.088 | 0.0695 | 0.07 | 0.07 | 0.06 |  |  | 89
| 1 | 6 | 0.07 | 0.11 | 0.09 | 0.07 | 0.1 | 0.1 | 0.09 | 0.088 | 0.0625 | 0.08 | 0.07 | 0.06
| Mean |  | 0.07 | 0.115 | 0.09 | 0.07 | 0.098333 | 0.1 | 0.083333 | 0.088 | 0.06845 | 0.076667 | 0.07 | 0.06
| Median |  | 0.07 | 0.105 | 0.09 | 0.07 | 0.1 | 0.1 | 0.085 | 0.088 | 0.0695 | 0.08 | 0.07 | 0.06
| Std Dev. |  | 0 | 0.037815 | 1.5202e-17 | 0 | 0.0040825 | 1.5202e-17 | 0.008165 | 0.0010954 | 0.0042094 | 0.005164 | 0 | 0
| Rel.Std.Dev. |  | 0% | 32.88% | 1.689e-14% | 0% | 4.152% | 1.52e-14% | 9.798% | 1.245% | 6.15% | 6.736% | 0% | 0%
| PDM3 |  | -13.83% | 41.57% | 10.79% | -13.83% | 21.05% | 23.1% | 2.587% | 8.332% | -15.74% | -5.62% | -13.83% | -26.14%
| Z-Score (Absolute) |  | 0.67 | 1.35 | 0.67 | 0.67 | 1.24 | 1.35 | 0.22 | 0.54 | 0.78 | 0.22 | 0.67 | 1.35
| Analytical results for La in OREAS 464 (Certified Value 1.06 wt.%) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | > 1 | > 1 | 0.485 | 0.497 | 0.434 | 1.2169 | 0.57946 | 1.1371 | > 0.5 | > 0.5 | 1.2141 | 1.1544 | 1.11 |  | 1
| 1 | 2 | 0.99 | > 1 | 0.473 | 0.494 | 0.436 | 1.2351 | 0.59222 | 1.1697 | > 0.5 | > 0.5 | 1.1776 | 1.1583 | 1.13 |  | 0
| 1 | 3 | 0.982 | > 1 | 0.478 | 0.499 | 0.442 | 1.2133 | 0.63512 | 1.1485 | > 0.5 | > 0.5 | 1.1872 | 1.2001 | 1.15 |  | 16
| 1 | 4 | > 1 | > 1 | 0.476 | 0.51 | 0.458 | 1.2224 | 0.57416 | 1.1307 | > 0.5 | > 0.5 | 1.1844 | 1.1442 | 1.14 |  | 1.0644
| 1 | 5 | 0.976 | > 1 | 0.474 | 0.522 | 0.435 | 1.2048 | 0.53198 | 1.1227 | > 0.5 | > 0.5 | 1.1948 | 1.1695 | 1.15 |  | 90
| 1 | 6 | 1 | > 1 | 0.478 | 0.51 | 0.451 | 1.2161 | 0.64198 | 1.1601 | > 0.5 | > 0.5 | 1.2408 | 1.2058 | 1.14
| Mean |  | 0.987 |  | 0.47733 | 0.50533 | 0.44267 | 1.2181 | 0.59249 | 1.1448 |  |  | 1.1998 | 1.172 | 1.1367
| Median |  | 0.986 |  | 0.477 | 0.5045 | 0.439 | 1.2165 | 0.58584 | 1.1428 |  |  | 1.191 | 1.1639 | 1.14
| Std Dev. |  | 0.010392 |  | 0.004274 | 0.010577 | 0.0098319 | 0.010148 | 0.04107 | 0.01797 |  |  | 0.023679 | 0.025342 | 0.015055
| Rel.Std.Dev. |  | 1.053% |  | 0.8954% | 2.093% | 2.221% | 0.8331% | 6.932% | 1.57% |  |  | 1.974% | 2.162% | 1.325%
| PDM3 |  | -7.273% |  | -55.16% | -52.53% | -58.41% | 14.44% | -44.34% | 7.553% |  |  | 12.72% | 10.11% | 6.787%
| Z-Score (Absolute) |  | 0.34 | NA | 2.68 | 2.55 | 2.84 | 0.72 | 2.15 | 0.38 | NA | NA | 0.63 | 0.51 | 0.34
| Analytical results for Li in OREAS 464 (Certified Value 8.26 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 2
2
| 1 | 1 | 6.8 | 6.9 | 7.6 | 8.3 | 5.1 | 8.583 | 7.3 | 9 | 7.9 | 3.1 | 9.852 | 9 | 10 | 8 | 1
| 1 | 2 | 6.9 | 6.7 | 7.6 | 8.1 | 4.5 | 8.3301 | 7.4 | 8 | 7.9 | 3 | 10.1666 | 10 | 10 | 8 | 16
| 1 | 3 | 7.2 | 6.2 | 7.9 | 8.2 | 4.7 | 7.3944 | 7.3 | 9 | 8 | 3 | 9.8194 | 9 | 10 | 8 | 16
| 1 | 4 | 7.2 | 6.6 | 7.7 | 8.3 | 4.6 | 7.441 | 7.3 | 9 | 7.8 | 3 | 10.1636 | 9 | 10 | 7.5 | 8.2553
| 1 | 5 | 7.3 | 7.2 | 7.7 | 8.3 | 4.3 | 8.0299 | 7.2 | 8 | 8.2 | 3.4 | 9.8651 | 9 | 10 | 8 | 91
| 1 | 6 | 7.5 | 6.8 | 7.8 | 8.3 | 4.1 | 8.4849 | 7.1 | 9 | 8 | 3.2 | 10.2536 | 10 | 10 | 8
| Mean |  | 7.15 | 6.7333 | 7.7167 | 8.25 | 4.55 | 8.0439 | 7.2667 | 8.6667 | 7.9667 | 3.1167 | 10.02 | 9.3333 | 10 | 7.9167
| Median |  | 7.2 | 6.75 | 7.7 | 8.3 | 4.55 | 8.18 | 7.3 | 9 | 7.95 | 3.05 | 10.0143 | 9 | 10 | 8
| Std Dev. |  | 0.25884 | 0.33267 | 0.1169 | 0.083666 | 0.34496 | 0.52011 | 0.10328 | 0.5164 | 0.13663 | 0.16021 | 0.1945 | 0.5164 | 0 | 0.20412
| Rel.Std.Dev. |  | 3.62% | 4.941% | 1.515% | 1.014% | 7.582% | 6.466% | 1.421% | 5.958% | 1.715% | 5.14% | 1.941% | 5.533% | 0% | 2.578%
| PDM3 |  | -13.39% | -18.44% | -6.525% | -0.06453% | -44.88% | -2.561% | -11.98% | 4.983% | -3.497% | -62.25% | 21.38% | 13.06% | 21.13% | -4.102%
| Z-Score (Absolute) |  | 0.7 | 1.07 | 0.2 | 0.27 | 3.02 | 0.09 | 0.6 | 0.64 | 0.02 | 4.34 | 1.85 | 1.24 | 1.83 | 0.02
| Analytical results for Lu in OREAS 464 (Certified Value 1.57 ppm) | [15 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 2.38 | 1.85 | 1.74 | 1.62 | 1.79 | 1.3708 | 1.2 | 1.41 | 1.05 | 1.38 | 2.3137 | 1.7 | 1.58 |  | 1
| 1 | 2 | 2.31 | 1.83 | 1.73 | 1.56 | 1.76 | 1.3515 | 1.2 | 1.49 | 1.13 | 1.4 | 2.2168 | 1.6 | 1.6 |  | 1
| 1 | 3 | 2.33 | 1.6 | 1.72 | 1.58 | 1.76 | 1.3216 | 1.2 | 1.48 | 1.19 | 1.38 | 2.2014 | 1.7 | 1.64 |  | 16
| 1 | 4 | 2.35 | 1.76 | 1.69 | 1.55 | 1.81 | 1.3831 | 1.2 | 1.49 | 1.07 | 1.4 | 2.1555 | 1.7 | 1.6 |  | 1.5722
| 1 | 5 | 2.32 | 1.67 | 1.82 | 1.61 | 1.84 | 1.2648 | 1.2 | 1.49 | 1.03 | 1.41 | 2.1625 | 1.8 | 1.66 |  | 92
| 1 | 6 | 2.33 | 1.83 | 1.73 | 1.58 | 1.78 | 1.2717 | 1.2 | 1.37 | 1.09 | 1.35 | 2.3384 | 1.6 | 1.64
| Mean |  | 2.3367 | 1.7567 | 1.7383 | 1.5833 | 1.79 | 1.3272 | 1.2 | 1.455 | 1.0933 | 1.3867 | 2.2314 | 1.6833 | 1.62
| Median |  | 2.33 | 1.795 | 1.73 | 1.58 | 1.785 | 1.3365 | 1.2 | 1.485 | 1.08 | 1.39 | 2.2091 | 1.7 | 1.62
| Std Dev. |  | 0.025033 | 0.10152 | 0.043551 | 0.027325 | 0.030984 | 0.050238 | 0 | 0.052058 | 0.058538 | 0.021602 | 0.077261 | 0.075277 | 0.030984
| Rel.Std.Dev. |  | 1.071% | 5.779% | 2.505% | 1.726% | 1.731% | 3.785% | 0% | 3.578% | 5.354% | 1.558% | 3.462% | 4.472% | 1.913%
| PDM3 |  | 48.63% | 11.74% | 10.57% | 0.7105% | 13.86% | -15.58% | -23.67% | -7.452% | -30.46% | -11.8% | 41.93% | 7.071% | 3.043%
| Z-Score (Absolute) |  | 2.84 | 0.54 | 0.4 | 0.15 | 0.67 | 1.16 | 1.67 | 0.59 | 2.09 | 0.93 | 2.43 | 0.25 | 0
| Analytical results for Mg in OREAS 464 (Certified Value 0.825 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.75 | 0.78 | 0.36 | 0.37 | 0.32 | 0.8939 | 0.78 | 0.85 | 0.863 | 0.897 | 0.77851 | 0.76 | 0.8455 | 0.85 | 1
| 1 | 2 | 0.75 | 0.77 | 0.35 | 0.37 | 0.33 | 0.84191 | 0.81 | 0.85 | 0.859 | 0.9153 | 0.77928 | 0.83 | 0.851 | 0.85 | 2
| 1 | 3 | 0.74 | 0.77 | 0.35 | 0.37 | 0.33 | 0.84424 | 0.81 | 0.85 | 0.864 | 0.8939 | 0.78226 | 0.8 | 0.839 | 0.89 | 16
| 1 | 4 | 0.77 | 0.77 | 0.35 | 0.38 | 0.34 | 0.86157 | 0.81 | 0.85 | 0.852 | 0.8702 | 0.79186 | 0.8 | 0.847 | 0.87 | 0.82488
| 1 | 5 | 0.74 | 0.75 | 0.35 | 0.39 | 0.33 | 0.85368 | 0.8 | 0.84 | 0.862 | 0.8596 | 0.78656 | 0.78 | 0.8495 | 0.85 | 93
| 1 | 6 | 0.76 | 0.79 | 0.35 | 0.38 | 0.33 | 0.87992 | 0.81 | 0.86 | 0.867 | 0.8556 | 0.78868 | 0.82 | 0.849 | 0.86
| Mean |  | 0.75167 | 0.77167 | 0.35167 | 0.37667 | 0.33 | 0.86254 | 0.80333 | 0.85 | 0.86117 | 0.88193 | 0.78453 | 0.79833 | 0.84683 | 0.86167
| Median |  | 0.75 | 0.77 | 0.35 | 0.375 | 0.33 | 0.85762 | 0.81 | 0.85 | 0.8625 | 0.88205 | 0.78441 | 0.8 | 0.848 | 0.855
| Std Dev. |  | 0.01169 | 0.013292 | 0.0040825 | 0.008165 | 0.0063246 | 0.02062 | 0.012111 | 0.0063246 | 0.0051929 | 0.023721 | 0.0053701 | 0.025626 | 0.0042973 | 0.016021
| Rel.Std.Dev. |  | 1.555% | 1.722% | 1.161% | 2.168% | 1.917% | 2.391% | 1.508% | 0.7441% | 0.603% | 2.69% | 0.6845% | 3.21% | 0.5075% | 1.859%
| PDM3 |  | -8.875% | -6.451% | -57.37% | -54.34% | -59.99% | 4.565% | -2.612% | 3.046% | 4.399% | 6.917% | -4.892% | -3.218% | 2.662% | 4.46%
| Z-Score (Absolute) |  | 0.61 | 0.36 | 5.53 | 5.22 | 5.8 | 0.76 | 0.03 | 0.61 | 0.74 | 1 | 0.2 | 0.03 | 0.57 | 0.75
| Analytical results for Mn in OREAS 464 (Certified Value 0.27 wt.%) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.239 | 0.267 | 0.0964 | 0.0976 | 0.086 | 0.27101 | 0.27551 | 0.2746 | 0.277 | 0.2903 | 0.242 | 0.2815 | 0.282 |  | 1
| 1 | 2 | 0.238 | 0.268 | 0.0933 | 0.0957 | 0.0851 | 0.26505 | 0.27892 | 0.2646 | 0.277 | 0.3021 | 0.2532 | 0.2815 | 0.279 |  | 3
| 1 | 3 | 0.236 | 0.269 | 0.0958 | 0.0969 | 0.0903 | 0.27094 | 0.2787 | 0.2667 | 0.279 | 0.2902 | 0.2322 | 0.2805 | 0.29 |  | 16
| 1 | 4 | 0.242 | 0.277 | 0.093 | 0.0972 | 0.0911 | 0.26964 | 0.2738 | 0.2713 | 0.272 | 0.2824 | 0.2479 | 0.281 | 0.286 |  | 0.26998
| 1 | 5 | 0.232 | 0.277 | 0.093 | 0.102 | 0.0854 | 0.26759 | 0.26937 | 0.2602 | 0.277 | 0.2798 | 0.2456 | 0.281 | 0.283 |  | 94
| 1 | 6 | 0.24 | 0.291 | 0.0939 | 0.102 | 0.0913 | 0.27426 | 0.2675 | 0.2756 | 0.28 | 0.2772 | 0.251 | 0.281 | 0.285
| Mean |  | 0.23783 | 0.27483 | 0.094233 | 0.098567 | 0.0882 | 0.26975 | 0.27397 | 0.26883 | 0.277 | 0.287 | 0.24532 | 0.28108 | 0.28417
| Median |  | 0.2385 | 0.273 | 0.0936 | 0.0974 | 0.08815 | 0.27029 | 0.27465 | 0.269 | 0.277 | 0.2863 | 0.24675 | 0.281 | 0.284
| Std Dev. |  | 0.0034881 | 0.0090866 | 0.0014949 | 0.002734 | 0.0029907 | 0.0031648 | 0.0047391 | 0.0060355 | 0.0027568 | 0.0091453 | 0.0075399 | 0.00037639 | 0.0037639
| Rel.Std.Dev. |  | 1.467% | 3.306% | 1.586% | 2.774% | 3.391% | 1.173% | 1.73% | 2.245% | 0.9952% | 3.187% | 3.074% | 0.1339% | 1.325%
| PDM3 |  | -11.91% | 1.798% | -65.1% | -63.49% | -67.33% | -0.08555% | 1.477% | -0.424% | 2.601% | 6.305% | -9.135% | 4.113% | 5.255%
| Z-Score (Absolute) |  | 1.49 | 0.24 | 8.21 | 8 | 8.49 | 0 | 0.2 | 0.04 | 0.34 | 0.81 | 1.14 | 0.53 | 0.67
| Analytical results for Mo in OREAS 464 (Certified Value 77 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
1
| 1 | 1 | 75.8 | 77.3 | 82.5 | 81.6 | 81.8 | 73.4024 | 75.7 | 81 | 80.7 | 77.7 | 73.5348 | 66.5 |  |  | 1
| 1 | 2 | 75.4 | 76.8 | 82.7 | 79.6 | 81.2 | 73.3387 | 77 | 83 | 79.4 | 73.8 | 74.8474 | 68 |  |  | 0
| 1 | 3 | 75.5 | 71.4 | 82 | 80 | 80.2 | 72.8595 | 75.6 | 82 | 80.9 | 77.2 | 73.4083 | 69 |  |  | 16
| 1 | 4 | 76.5 | 73.8 | 85.3 | 81.6 | 83.2 | 73.8951 | 75.4 | 79 | 79.7 | 73.9 | 73.8087 | 68.5 |  |  | 77.135
| 1 | 5 | 74.6 | 79.8 | 84.6 | 84 | 84.1 | 73.2376 | 73.4 | 78 | 80.3 | 77.9 | 73.1467 | 69.5 |  |  | 95
| 1 | 6 | 76.9 | 76.7 | 81.1 | 81.5 | 81.5 | 73.7259 | 73.5 | 80 | 80.1 | 76.1 | 73.2136 | 69.5
| Mean |  | 75.7833 | 75.9667 | 83.0333 | 81.3833 | 82 | 73.4099 | 75.1 | 80.5 | 80.1833 | 76.1 | 73.6599 | 68.5
| Median |  | 75.65 | 76.75 | 82.6 | 81.55 | 81.65 | 73.3706 | 75.5 | 80.5 | 80.2 | 76.65 | 73.4716 | 68.75
| Std Dev. |  | 0.82321 | 2.9412 | 1.5996 | 1.5523 | 1.4156 | 0.3669 | 1.3971 | 1.8708 | 0.57417 | 1.8515 | 0.62828 | 1.1402
| Rel.Std.Dev. |  | 1.086% | 3.872% | 1.926% | 1.907% | 1.726% | 0.4998% | 1.86% | 2.324% | 0.7161% | 2.433% | 0.8529% | 1.664%
| PDM3 |  | -1.752% | -1.515% | 7.647% | 5.508% | 6.307% | -4.829% | -2.638% | 4.363% | 3.952% | -1.342% | -4.505% | -11.19%
| Z-Score (Absolute) |  | 0.05 | 0.01 | 1.39 | 1.07 | 1.19 | 0.52 | 0.19 | 0.89 | 0.83 | 0.01 | 0.47 | 1.5
| Analytical results for Na in OREAS 464 (Certified Value 0.149 wt.%) | [15 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.13 | 0.14 | 0.14 | 0.16 | 0.14 | 0.14 | 0.15 | 0.158 | 0.142 | 0.15008 | 0.17 | 0.14 | 0.16 |  | 1
| 1 | 2 | 0.13 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.15 | 0.157 | 0.1452 | 0.15615 | 0.17 | 0.14 | 0.15 |  | 4
| 1 | 3 | 0.13 | 0.16 | 0.14 | 0.16 | 0.13 | 0.15 | 0.14 | 0.159 | 0.1384 | 0.15354 | 0.17 | 0.14 | 0.16 |  | 16
| 1 | 4 | 0.13 | 0.16 | 0.14 | 0.16 | 0.14 | 0.15 | 0.15 | 0.155 | 0.1361 | 0.15589 | 0.17 | 0.14 | 0.16 |  | 0.14862
| 1 | 5 | 0.13 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.14 | 0.158 | 0.1357 | 0.15531 | 0.18 | 0.14 | 0.15 |  | 96
| 1 | 6 | 0.14 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.14 | 0.159 | 0.1329 | 0.15521 | 0.17 | 0.14 | 0.15
| Mean |  | 0.13167 | 0.15167 | 0.14 | 0.16 | 0.13833 | 0.14833 | 0.145 | 0.15767 | 0.13838 | 0.15436 | 0.17167 | 0.14 | 0.155
| Median |  | 0.13 | 0.15 | 0.14 | 0.16 | 0.14 | 0.15 | 0.145 | 0.158 | 0.13725 | 0.15526 | 0.17 | 0.14 | 0.155
| Std Dev. |  | 0.0040825 | 0.0075277 | 0 | 0 | 0.0040825 | 0.0040825 | 0.0054772 | 0.0015055 | 0.0045173 | 0.0022864 | 0.0040825 | 0 | 0.0054772
| Rel.Std.Dev. |  | 3.101% | 4.963% | 0% | 0% | 2.951% | 2.752% | 3.777% | 0.9549% | 3.264% | 1.481% | 2.378% | 0% | 3.534%
| PDM3 |  | -11.41% | 2.049% | -5.801% | 7.656% | -6.922% | -0.1939% | -2.437% | 6.086% | -6.889% | 3.863% | 15.51% | -5.801% | 4.292%
| Z-Score (Absolute) |  | 1.35 | 0.27 | 0.67 | 0.94 | 0.81 | 0 | 0.27 | 0.76 | 0.81 | 0.49 | 1.89 | 0.67 | 0.54
| Analytical results for Nb in OREAS 464 (Indicative Value 1598 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | > 500 | > 500 | > 500 | > 500 | > 500 | 1622.7466 | 1633.8 | 1356 | > 500 | 1860 |  |  |  |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | > 500 | > 500 | 1679.193 | 1653.4 | 1063 | > 500 | 1880 |  |  |  |  | 5
| 1 | 3 | > 500 | > 500 | > 500 | > 500 | > 500 | 1684.3296 | 1629.3 | 1167 | > 500 | 1940 |  |  |  |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 1643.3773 | 1646.1 | 1224 | > 500 | 1970 |  |  |  |  | 1597.7592
| 1 | 5 | > 500 | > 500 | > 500 | > 500 | > 500 | 1630.7065 | 1612.3 | 939 | > 500 | 1950 |  |  |  |  | 27
| 1 | 6 | > 500 | > 500 | > 500 | > 500 | > 500 | 1693.5686 | 1598.4 | 1300 | > 500 | 1970
| Mean |  | 07: 1658.987; 08: 1628.8833; 09: 1174.8333; 22: 1928.3333
| Median |  | 07: 1661.2852; 08: 1631.55; 09: 1195.5; 22: 1945
| Std Dev. |  | 07: 30.3421; 08: 20.6207; 09: 154.3314; 22: 47.0815
| Rel.Std.Dev. |  | 07: 1.829%; 08: 1.266%; 09: 13.14%; 22: 2.442%
| PDM3 |  | 07: 3.832%; 08: 1.948%; 09: -26.47%; 22: 20.69%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.07 | 0.07 | 2.11 | NA | 1.28
| Analytical results for Nd in OREAS 464 (Certified Value 0.959 wt.%) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | > 1 | > 1 | 0.951 | > 1 | 0.941 | 1.0498 | 0.60099 | 0.9651 | > 0.5 | 0.92443 | 0.92704 | 0.97225 | 0.947 |  | 1
| 1 | 2 | 0.991 | > 1 | 0.946 | > 1 | 0.932 | 1.0482 | 0.65359 | 0.9773 | > 0.5 | 0.92877 | 0.91585 | 0.9743 | 0.947 |  | 0
| 1 | 3 | 1 | 0.918 | 0.966 | > 1 | 0.924 | 1.0415 | 0.65278 | 0.9738 | > 0.5 | 0.97265 | 0.91595 | 1.0126 | 0.964 |  | 16
| 1 | 4 | 0.994 | 0.975 | 0.959 | > 1 | 0.951 | 1.0411 | 0.60996 | 0.9752 | > 0.5 | 0.93701 | 0.92544 | 0.967 | 0.97 |  | 0.95892
| 1 | 5 | 0.983 | 0.968 | 0.944 | > 1 | 0.983 | 1.0471 | 0.57539 | 0.9632 | > 0.5 | 0.95021 | 0.92112 | 0.9827 | 0.968 |  | 97
| 1 | 6 | 0.985 | 1 | 0.938 | > 1 | 0.954 | 1.0506 | 0.64421 | 0.9628 | > 0.5 | 0.92692 | 0.9271 | 1.009 | 0.954
| Mean |  | 0.9906 | 0.96525 | 0.95067 |  | 0.9475 | 1.0464 | 0.62282 | 0.96957 |  | 0.94 | 0.92208 | 0.98631 | 0.95833
| Median |  | 0.991 | 0.9715 | 0.9485 |  | 0.946 | 1.0477 | 0.62709 | 0.96945 |  | 0.93289 | 0.92328 | 0.9785 | 0.959
| Std Dev. |  | 0.0068775 | 0.034364 | 0.010309 |  | 0.020734 | 0.0041104 | 0.032229 | 0.0065686 |  | 0.018542 | 0.0052625 | 0.019666 | 0.010367
| Rel.Std.Dev. |  | 0.6943% | 3.56% | 1.084% |  | 2.188% | 0.3928% | 5.175% | 0.6775% |  | 1.973% | 0.5707% | 1.994% | 1.082%
| PDM3 |  | 3.303% | 0.6598% | -0.861% |  | -1.191% | 9.122% | -35.05% | 1.11% |  | -1.974% | -3.842% | 2.856% | -0.06148%
| Z-Score (Absolute) |  | 1.19 | 0.25 | 0.28 | NA | 0.4 | 3.24 | 12.34 | 0.41 | NA | 0.67 | 1.33 | 1.03 | 0
| Analytical results for Ni in OREAS 464 (Certified Value 89 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
1
| 1 | 1 | 362 | 85.7 | 84.5 | 88.1 | 83.4 | 83.804 | 95 | 92.8 | 95 | 80 | 100 |  |  |  | 1
| 1 | 2 | 357 | 84.5 | 83.6 | 84.6 | 82 | 85.0529 | 94 | 91.3 | 96 | 80 | 108 |  |  |  | 6
| 1 | 3 | 357 | 78.1 | 86.1 | 86.7 | 79.6 | 84.0586 | 89 | 92.9 | 92 | 85 | 100 |  |  |  | 16
| 1 | 4 | 365 | 83.9 | 84.6 | 85.4 | 83 | 84.6677 | 99 | 92.8 | 92 | 80 | 104 |  |  |  | 88.5748
| 1 | 5 | 358 | 86.2 | 83.8 | 89.7 | 84.8 | 85.676 | 89 | 91 | 90 | 85 | 102 |  |  |  | 98
| 1 | 6 | 369 | 85.7 | 84.4 | 86.6 | 87.5 | 84.8258 | 87 | 92 | 88 | 80 | 104
| Mean |  | 361.3333 | 84.0167 | 84.5 | 86.85 | 83.3833 | 84.6809 | 92.1667 | 92.1333 | 92.1667 | 81.6667 | 103
| Median |  | 360 | 85.1 | 84.45 | 86.65 | 83.2 | 84.7468 | 91.5 | 92.4 | 92 | 80 | 103
| Std Dev. |  | 4.9261 | 3.0229 | 0.88091 | 1.8404 | 2.6581 | 0.67912 | 4.5789 | 0.83347 | 2.9944 | 2.582 | 3.0332
| Rel.Std.Dev. |  | 1.363% | 3.598% | 1.042% | 2.119% | 3.188% | 0.802% | 4.968% | 0.9046% | 3.249% | 3.162% | 2.945%
| PDM3 |  | 307.9% | -5.146% | -4.6% | -1.947% | -5.861% | -4.396% | 4.055% | 4.018% | 4.055% | -7.799% | 16.29%
| Z-Score (Absolute) |  | 35.71 | 0.21 | 0.31 | 0 | 0.45 | 0.28 | 0.69 | 0.69 | 0.69 | 0.67 | 2.1
| Analytical results for P in OREAS 464 (Certified Value 1.36 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | > 1 | > 1 | 0.614 | 0.654 | 0.567 | 1.297 | 1.45 | 1.37 | 1.4155 | 1.2593 | 1.3 | 1.32 | 1.35 |  | 1
| 1 | 2 | > 1 | > 1 | 0.598 | 0.646 | 0.569 | 1.319 | 1.45 | 1.36 | 1.4482 | 1.2714 | 1.3 | 1.3185 | 1.49 |  | 7
| 1 | 3 | > 1 | > 1 | 0.605 | 0.655 | 0.584 | 1.321 | 1.46 | 1.36 | 1.395 | 1.2731 | 1.3 | 1.334 | 1.59 |  | 16
| 1 | 4 | > 1 | > 1 | 0.602 | 0.663 | 0.593 | 1.304 | 1.42 | 1.33 | 1.383 | 1.2943 | 1.3 | 1.321 | 1.46 |  | 1.3595
| 1 | 5 | > 1 | > 1 | 0.601 | 0.678 | 0.568 | 1.313 | 1.42 | 1.36 | 1.362 | 1.2873 | 1.3 | 1.306 | 1.42 |  | 99
| 1 | 6 | > 1 | > 1 | 0.605 | 0.672 | 0.586 | 1.311 | 1.48 | 1.37 | 1.3581 | 1.2981 | 1.4 | 1.3175 | 1.53
| Mean |  |  |  | 0.60417 | 0.66133 | 0.57783 | 1.3108 | 1.4467 | 1.3583 | 1.3936 | 1.2806 | 1.3167 | 1.3195 | 1.4733
| Median |  |  |  | 0.6035 | 0.659 | 0.5765 | 1.312 | 1.45 | 1.36 | 1.389 | 1.2802 | 1.3 | 1.3193 | 1.475
| Std Dev. |  |  |  | 0.0054924 | 0.012028 | 0.011197 | 0.0090866 | 0.023381 | 0.01472 | 0.034158 | 0.01505 | 0.040825 | 0.0089443 | 0.084063
| Rel.Std.Dev. |  |  |  | 0.9091% | 1.819% | 1.938% | 0.6932% | 1.616% | 1.084% | 2.451% | 1.175% | 3.101% | 0.6779% | 5.706%
| PDM3 |  |  |  | -55.56% | -51.36% | -57.5% | -3.582% | 6.41% | -0.08781% | 2.509% | -5.807% | -3.153% | -2.944% | 8.371%
| Z-Score (Absolute) |  | NA | NA | 6.24 | 5.74 | 6.47 | 0.05 | 1.14 | 0.37 | 0.67 | 0.32 | 0 | 0.02 | 1.17
| Analytical results for Pb in OREAS 464 (Certified Value 205 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 206 | 219 | 215 | 212 | 205 | 186.8226 | 175 | 190 | 146.2 | 255.535 | 260 | 173 |  |  | 1
| 1 | 2 | 199 | 211 | 226 | 227 | 210 | 183.1645 | 181 | 191 | 144.5 | 255.815 | 260 | 171 |  |  | 8
| 1 | 3 | 200 | 227 | 216 | 217 | 203 | 182.5035 | 177 | 199 | 152.2 | 255.201 | 260 | 174 |  |  | 16
| 1 | 4 | 205 | 208 | 229 | 231 | 207 | 185.4246 | 181 | 188 | 151.2 | 255.742 | 270 | 174 |  |  | 204.9508
| 1 | 5 | 199.5 | 206 | 230 | 220 | 213 | 186.6728 | 178 | 193 | 151.7 | 254.587 | 260 | 176 |  |  | 100
| 1 | 6 | 209 | 217 | 218 | 210 | 203 | 189.5934 | 173 | 194 | 145.6 | 256.494 | 250 | 171
| Mean |  | 203.0833 | 214.6667 | 222.3333 | 219.5 | 206.8333 | 185.6969 | 177.5 | 192.5 | 148.5667 | 255.5623 | 260 | 173.1667
| Median |  | 202.5 | 214 | 222 | 218.5 | 206 | 186.0487 | 177.5 | 192 | 148.7 | 255.6385 | 260 | 173.5
| Std Dev. |  | 4.1523 | 7.8655 | 6.7725 | 8.2644 | 4.0208 | 2.611 | 3.2094 | 3.8341 | 3.4898 | 0.6398 | 6.3246 | 1.9408
| Rel.Std.Dev. |  | 2.045% | 3.664% | 3.046% | 3.765% | 1.944% | 1.406% | 1.808% | 1.992% | 2.349% | 0.2503% | 2.433% | 1.121%
| PDM3 |  | -0.9112% | 4.741% | 8.481% | 7.099% | 0.9185% | -9.394% | -13.39% | -6.075% | -27.51% | 24.69% | 26.86% | -15.51%
| Z-Score (Absolute) |  | 0.07 | 0.36 | 0.64 | 0.54 | 0.07 | 0.71 | 1.01 | 0.46 | 2.08 | 1.86 | 2.03 | 1.17
| Analytical results for Pr in OREAS 464 (Certified Value 2601 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2766.0865 | 1499.4 | 2526.18 | 993 | 2614.19 | 2738.0649 | 2513.1 | 2460 |  | 1
| 1 | 2 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2705.2712 | 1631.1 | 2592.95 | 997 | 2677.05 | 2640.5558 | 2535.7 | 2470 |  | 0
| 1 | 3 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2603.8712 | 1628.9 | 2574.9 | 1050 | 2601.65 | 2604.5449 | 2621.6 | 2490 |  | 16
| 1 | 4 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2659.473 | 1517.4 | 2551.77 | 1000 | 2615.49 | 2670.0232 | 2514.4 | 2490 |  | 2600.6204
| 1 | 5 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2714.0279 | 1421.4 | 2519.2 | 1040 | 2702.63 | 2691.2284 | 2540.8 | 2500 |  | 101
| 1 | 6 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 2641.0841 | 1631.5 | 2561.15 | 1060 | 2592.96 | 2786.2845 | 2636.1 | 2500
| Mean | [6 empty] | 2681.6357 | 1554.95 | 2554.3583 | 1023.3333 | 2633.995 | 2688.4503 | 2560.2833 | 2485
| Median | [6 empty] | 2682.3721 | 1573.15 | 2556.46 | 1020 | 2614.84 | 2680.6258 | 2538.25 | 2490
| Std Dev. | [6 empty] | 58.1769 | 88.8355 | 28.2639 | 29.9711 | 44.7867 | 65.9006 | 54.4489 | 16.4317
| Rel.Std.Dev. | [6 empty] | 2.169% | 5.713% | 1.106% | 2.929% | 1.7% | 2.451% | 2.127% | 0.6612%
| PDM3 | [6 empty] | 3.115% | -40.21% | -1.779% | -60.65% | 1.283% | 3.377% | -1.551% | -4.446%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.83 | 6.73 | 0.02 | 10.29 | 0.51 | 0.88 | 0.02 | 0.49
| Analytical results for Rb in OREAS 464 (Certified Value 4.22 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.7 | 4.5 | 4.3 | 4.1 | 3.9 | 4.148 | 4 | 4.7 | 3.46 | 4.17 | 5.0242 | 4 | 4.6 |  | 1
| 1 | 2 | 4.8 | 4.6 | 4.5 | 4.1 | 3.9 | 4.0603 | 4.1 | 4.9 | 3.29 | 4 | 4.86 | 4 | 4.6 |  | 0
| 1 | 3 | 4.4 | 4.1 | 4.3 | 4 | 3.9 | 4.2091 | 4 | 4.7 | 3.42 | 4.06 | 5.2089 | 4 | 4.6 |  | 16
| 1 | 4 | 4.9 | 4.2 | 4.4 | 4.1 | 3.9 | 4.386 | 4 | 4.9 | 3.32 | 4.17 | 5.3312 | 4 | 4.4 |  | 4.2208
| 1 | 5 | 4.8 | 4.7 | 4.5 | 4.2 | 4.1 | 4.0062 | 3.9 | 4.6 | 3.47 | 4.39 | 5.0678 | 4 | 4.6 |  | 102
| 1 | 6 | 4.9 | 4.4 | 4.2 | 4.1 | 3.9 | 3.8792 | 4 | 4.8 | 3.43 | 4.21 | 5.0192 | 4 | 4.6
| Mean |  | 4.75 | 4.4167 | 4.3667 | 4.1 | 3.9333 | 4.1148 | 4 | 4.7667 | 3.3983 | 4.1667 | 5.0852 | 4 | 4.5667
| Median |  | 4.8 | 4.45 | 4.35 | 4.1 | 3.9 | 4.1041 | 4 | 4.75 | 3.425 | 4.17 | 5.046 | 4 | 4.6
| Std Dev. |  | 0.18708 | 0.23166 | 0.12111 | 0.063246 | 0.08165 | 0.1754 | 0.063246 | 0.12111 | 0.075211 | 0.13486 | 0.16426 | 0 | 0.08165
| Rel.Std.Dev. |  | 3.939% | 5.245% | 2.773% | 1.543% | 2.076% | 4.263% | 1.581% | 2.541% | 2.213% | 3.237% | 3.23% | 0% | 1.788%
| PDM3 |  | 12.54% | 4.64% | 3.455% | -2.862% | -6.811% | -2.511% | -5.232% | 12.93% | -19.49% | -1.283% | 20.48% | -5.232% | 8.194%
| Z-Score (Absolute) |  | 1.89 | 0.72 | 0.58 | 0.19 | 0.67 | 0.15 | 0.48 | 1.73 | 2.22 | 0 | 2.65 | 0.48 | 1.16
| Analytical results for Re in OREAS 464 (Certified Value < 0.03 ppm) | [15 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.049 | 0.005 | 0.017 | 0.012 | 0.02 | < 0.002 | 0.02 | 0.005 | 0.008 | < 0.05 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.051 | 0.005 | 0.017 | 0.01 | 0.021 | < 0.002 | < 0.02 | 0.003 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 18
| 1 | 3 | 0.051 | 0.004 | 0.017 | 0.012 | 0.021 | < 0.002 | < 0.02 | 0.004 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.051 | 0.008 | 0.017 | 0.014 | 0.02 | < 0.002 | < 0.02 | 0.004 | 0.003 | < 0.05 | < 0.1 |  |  |  | < 0.03
| 1 | 5 | 0.048 | 0.002 | 0.013 | 0.017 | 0.021 | < 0.002 | < 0.02 | 0.004 | 0.003 | < 0.05 | < 0.1 |  |  |  | 103
| 1 | 6 | 0.047 | 0.015 | 0.018 | 0.011 | 0.02 | < 0.002 | < 0.02 | 0.004 | 0.003 | < 0.05 | < 0.1
| Mean |  | 0.0495 | 0.0065 | 0.0165 | 0.012667 | 0.0205 |  | 0.02 | 0.004 | 0.00425
| Median |  | 0.05 | 0.005 | 0.017 | 0.012 | 0.0205 |  | 0.02 | 0.004 | 0.003
| Std Dev. |  | 01: 0.0017607; 02: 0.0045935; 03: 0.0017607; 05: 0.0025033; 06: 0.00054772; 09: 0.00063246; 10: 0.0025
| Rel.Std.Dev. |  | 01: 3.557%; 02: 70.67%; 03: 10.67%; 05: 19.76%; 06: 2.672%; 09: 15.81%; 10: 58.82%
PDM3
| Z-Score (Absolute) |  | 3.33 | 0.81 | 0.22 | 0.22 | 0.54 | NA | 0.49 | 1.05 | 1.15 | NA | NA
| Analytical results for S in OREAS 464 (Certified Value 983 ppm) | [15 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 960 | 963 | 1100 | 1050 | 1000 |  |  | 1
| 1 | 2 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 920 | 1095 | 1100 | 1050 | 1000 |  |  | 11
| 1 | 3 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 940 | 934 | 1100 | 1050 | 1050 |  |  | 16
| 1 | 4 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 920 | 949 | 1200 | 1050 | 1000 |  |  | 983.4111
| 1 | 5 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 940 | 1277 | 1100 | 1050 | 1050 |  |  | 104
| 1 | 6 | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 930 | 972 | 1100 | 1050 | 1000
| Mean |  | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 935 | 1031.6667 | 1116.6667 | 1050 | 1016.6667
| Median |  | 900 | 1100 | 900 | 1000 | 1000 | 900 | 900 | 935 | 967.5 | 1100 | 1050 | 1000
| Std Dev. |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.1658 | 133.3022 | 40.8248 | 0 | 25.8199
| Rel.Std.Dev. |  | 0% | 0% | 0% | 0% | 0% | 0% | 0% | 1.622% | 12.92% | 3.656% | 0% | 2.54%
| PDM3 |  | -8.482% | 11.86% | -8.482% | 1.687% | 1.687% | -8.482% | -8.482% | -4.923% | 4.907% | 13.55% | 6.771% | 3.382%
| Z-Score (Absolute) |  | 0.82 | 0.98 | 0.82 | 0.08 | 0.08 | 0.82 | 0.82 | 0.51 | 0.08 | 1.13 | 0.53 | 0.23
| Analytical results for Sb in OREAS 464 (Certified Value 1.87 ppm) | [15 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
3
| 1 | 1 | 1.77 | 1.92 | 2.07 | 1.93 | 1.86 | 1.617 | 1.9 | 1.8 | 1.75 | 1.47 | 1.8 |  |  |  | 1
| 1 | 2 | 1.79 | 1.86 | 1.97 | 1.87 | 1.83 | 1.6035 | 1.9 | 1.8 | 1.74 | 1.39 | 1.8 |  |  |  | 0
| 1 | 3 | 1.83 | 1.69 | 2.02 | 1.94 | 1.83 | 1.807 | 1.9 | 2.1 | 1.8 | 1.46 | 1.8 |  |  |  | 16
| 1 | 4 | 1.83 | 1.87 | 2.06 | 1.95 | 1.87 | 1.6575 | 1.9 | 1.9 | 1.82 | 1.47 | 1.8 |  |  |  | 1.868
| 1 | 5 | 1.8 | 1.99 | 2.17 | 2.03 | 1.94 | 1.7175 | 1.9 | 1.9 | 1.83 | 1.49 | 1.6 |  |  |  | 105
| 1 | 6 | 1.98 | 1.9 | 2 | 1.94 | 1.86 | 1.7611 | 1.8 | 2 | 1.79 | 1.43 | 1.8
| Mean |  | 1.8333 | 1.8717 | 2.0483 | 1.9433 | 1.865 | 1.6939 | 1.8833 | 1.9167 | 1.7883 | 1.4517 | 1.7667
| Median |  | 1.815 | 1.885 | 2.04 | 1.94 | 1.86 | 1.6875 | 1.9 | 1.9 | 1.795 | 1.465 | 1.8
| Std Dev. |  | 0.075542 | 0.10028 | 0.070261 | 0.051251 | 0.040373 | 0.081614 | 0.040825 | 0.1169 | 0.03656 | 0.036009 | 0.08165
| Rel.Std.Dev. |  | 4.12% | 5.358% | 3.43% | 2.637% | 2.165% | 4.818% | 2.168% | 6.099% | 2.044% | 2.481% | 4.622%
| PDM3 |  | -1.857% | 0.1949% | 9.652% | 4.031% | -0.162% | -9.32% | 0.8195% | 2.604% | -4.266% | -22.29% | -5.426%
| Z-Score (Absolute) |  | 0.33 | 0.45 | 1.9 | 0.81 | 0 | 1.77 | 0.19 | 0.54 | 0.8 | 4.29 | 0.67
| Analytical results for Sc in OREAS 464 (Certified Value 141 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 134.5 | 147 | 141.5 | 150 | 143 | 132.49 | 120 | 140 | 83.6 | 149.9738 | 140 | 143 |  |  | 1
| 1 | 2 | 135 | 141.5 | 142 | 146.5 | 140.5 | 136.0164 | 126 | 135 | 84 | 151.1792 | 139 | 143 |  |  | 19
| 1 | 3 | 134 | 132 | 147 | 148.5 | 137 | 130.3956 | 126 | 138 | 85.8 | 148.6975 | 139 | 150 |  |  | 16
| 1 | 4 | 136.5 | 140 | 143.5 | 150.5 | 142 | 136.9 | 118 | 138 | 89 | 158.299 | 139 | 147 |  |  | 141.3799
| 1 | 5 | 133 | 134.5 | 144 | 152.5 | 146 | 134.9671 | 113 | 134 | 89 | 148.4705 | 139 | 146 |  |  | 106
| 1 | 6 | 137 | 141 | 142 | 148.5 | 138.5 | 130.8271 | 121 | 141 | 86 | 148.9277 | 139 | 145
| Mean |  | 135 | 139.3333 | 143.3333 | 149.4167 | 141.1667 | 133.5994 | 120.6667 | 137.6667 | 86.2333 | 150.9246 | 139.1667 | 145.6667
| Median |  | 134.75 | 140.5 | 142.75 | 149.25 | 141.25 | 133.7286 | 120.5 | 138 | 85.9 | 149.4507 | 139 | 145.5
| Std Dev. |  | 1.5166 | 5.3635 | 2.0412 | 2.0595 | 3.2352 | 2.7491 | 4.9666 | 2.7325 | 2.3441 | 3.7502 | 0.40825 | 2.6583
| Rel.Std.Dev. |  | 1.123% | 3.849% | 1.424% | 1.378% | 2.292% | 2.058% | 4.116% | 1.985% | 2.718% | 2.485% | 0.2934% | 1.825%
| PDM3 |  | -4.513% | -1.448% | 1.382% | 5.685% | -0.1508% | -5.503% | -14.65% | -2.626% | -39.01% | 6.751% | -1.565% | 3.032%
| Z-Score (Absolute) |  | 0.58 | 0.01 | 0.56 | 1.38 | 0.26 | 0.77 | 2.53 | 0.22 | 7.22 | 1.39 | 0.01 | 0.87
| Analytical results for Se in OREAS 464 (Certified Value < 50 ppm) | [15 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 13 | < 50 | 16 | 27 | 40 | 10.1443 | 28 | 35 | < 0.5 | 5.7 | 30 |  |  |  | 1
| 1 | 2 | 13 | < 50 | 18 | 27 | 40 | 13.5092 | 28 | 38 | < 0.5 | 5.5 | 35 |  |  |  | 15
| 1 | 3 | 13 | < 50 | 18 | 27 | 39 | 11.986 | 29 | 38 | < 0.5 | 1.2 | 30 |  |  |  | 16
| 1 | 4 | 14 | < 50 | 17 | 26 | 39 | 12.8004 | 29 | 36 | < 0.5 | 6.4 | 30 |  |  |  | < 50
| 1 | 5 | 12 | < 50 | 17 | 27 | 41 | 9.7127 | 27 | 35 | < 0.5 | 4.5 | 35 |  |  |  | 107
| 1 | 6 | 13 | < 50 | 16 | 26 | 39 | 13.1495 | 29 | 38 | < 0.5 | 3.6 | 30
| Mean |  | 13 |  | 17 | 26.6667 | 39.6667 | 11.8837 | 28.3333 | 36.6667 |  | 4.4833 | 31.6667
| Median |  | 13 |  | 17 | 27 | 39.5 | 12.3932 | 28.5 | 37 |  | 5 | 30
| Std Dev. |  | 0.63246 |  | 0.89443 | 0.5164 | 0.8165 | 1.6021 | 0.8165 | 1.5055 |  | 1.8841 | 2.582
| Rel.Std.Dev. |  | 4.865% |  | 5.261% | 1.936% | 2.058% | 13.48% | 2.882% | 4.106% |  | 42.02% | 8.154%
PDM3
| Z-Score (Absolute) |  | 0.92 | NA | 0.65 | 0 | 0.88 | 1 | 0.11 | 0.67 | NA | 1.5 | 0.34
| Analytical results for Sm in OREAS 464 (Certified Value 1485 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1603.3667 | 1019.4 | 1470.4 | 748 | 1446.62 | 1425.2329 | 1449.8 | 1490 |  | 1
| 1 | 2 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1569.56 | 1097.5 | 1494.5 | 745 | 1471.25 | 1396.2436 | 1457.6 | 1480 |  | 0
| 1 | 3 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1493.8714 | 1089.6 | 1491.4 | 712 | 1496.2 | 1451.6107 | 1519.9 | 1520 |  | 16
| 1 | 4 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1534.734 | 1048.3 | 1474.2 | 715 | 1454.61 | 1403.0298 | 1444.5 | 1510 |  | 1484.8664
| 1 | 5 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1608.871 | 987.3 | 1476.6 | 741 | 1511.9 | 1408.1493 | 1461.7 | 1520 |  | 108
| 1 | 6 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1531.924 | 1081.6 | 1485 | 721 | 1430.31 | 1444.2079 | 1507.9 | 1520
| Mean | [6 empty] | 1557.0545 | 1053.95 | 1482.0167 | 730.3333 | 1468.4817 | 1421.4124 | 1473.5667 | 1506.6667
| Median | [6 empty] | 1552.147 | 1064.95 | 1480.8 | 731 | 1462.93 | 1416.6911 | 1459.65 | 1515
| Std Dev. | [6 empty] | 44.9622 | 43.8176 | 9.7792 | 16.1204 | 30.9469 | 22.7749 | 32.0345 | 17.5119
| Rel.Std.Dev. | [6 empty] | 2.888% | 4.157% | 0.6599% | 2.207% | 2.107% | 1.602% | 2.174% | 1.162%
| PDM3 | [6 empty] | 4.862% | -29.02% | -0.1919% | -50.81% | -1.103% | -4.273% | -0.761% | 1.468%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 1.36 | 6.6 | 0.17 | 11.72 | 0.04 | 0.78 | 0.04 | 0.56
| Analytical results for Sn in OREAS 464 (Certified Value 35.7 ppm) | [15 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
2
| 1 | 1 | 33.8 | 39.2 | 42 | 37.4 | 37.2 | 31.714 | 36.3 | 37 | 28.6 | 37.5 | 37 |  |  |  | 1
| 1 | 2 | 33.5 | 37.9 | 40.3 | 36.4 | 36.9 | 33.4239 | 36.4 | 33 | 29 | 34.8 | 37 |  |  |  | 10
| 1 | 3 | 34.4 | 35.1 | 41.5 | 36.5 | 36.7 | 31.4885 | 36.6 | 33 | 29.2 | 36 | 36 |  |  |  | 16
| 1 | 4 | 33.6 | 37.1 | 42.2 | 37.4 | 37.9 | 32.9356 | 36.3 | 35 | 28.5 | 36.4 | 36 |  |  |  | 35.6779
| 1 | 5 | 33.7 | 40.5 | 41.3 | 38.2 | 38.7 | 32.8752 | 35.8 | 32 | 29.9 | 35.7 | 37 |  |  |  | 109
| 1 | 6 | 34.3 | 39.1 | 39.9 | 36.9 | 37.9 | 32.3009 | 35.6 | 35 | 29.6 | 36.3 | 36
| Mean |  | 33.8833 | 38.15 | 41.2 | 37.1333 | 37.55 | 32.4564 | 36.1667 | 34.1667 | 29.1333 | 36.1167 | 36.5
| Median |  | 33.75 | 38.5 | 41.4 | 37.15 | 37.55 | 32.5881 | 36.3 | 34 | 29.1 | 36.15 | 36.5
| Std Dev. |  | 0.37639 | 1.8971 | 0.92087 | 0.67429 | 0.75299 | 0.75545 | 0.38297 | 1.8348 | 0.55015 | 0.88863 | 0.54772
| Rel.Std.Dev. |  | 1.111% | 4.973% | 2.235% | 1.816% | 2.005% | 2.328% | 1.059% | 5.37% | 1.888% | 2.46% | 1.501%
| PDM3 |  | -5.03% | 6.929% | 15.48% | 4.08% | 5.247% | -9.029% | 1.37% | -4.236% | -18.34% | 1.23% | 2.304%
| Z-Score (Absolute) |  | 0.78 | 0.67 | 1.71 | 0.33 | 0.47 | 1.26 | 0 | 0.68 | 2.39 | 0.02 | 0.11
| Analytical results for Sr in OREAS 464 (Certified Value 2043 ppm) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 2290 | 2200 | 1070 | 1150 | 951 | 2093 | 1890.7 | 2194.9 | 1820 | 2022.37 | 2048.6222 | 1949 | 2065 | 1820 | 1
| 1 | 2 | 2310 | 2160 | 1050 | 1130 | 960 | 2072.5164 | 1912.8 | 2215.3 | 1840 | 1987.56 | 2094.1348 | 1953 | 2070 | 1830 | 1
| 1 | 3 | 2300 | 2160 | 1060 | 1135 | 966 | 2021 | 1894.3 | 2228.5 | 1860 | 2075.14 | 2068.5169 | 1814 | 2070 | 1860 | 16
| 1 | 4 | 2320 | 2130 | 1060 | 1170 | 1010 | 2052.2607 | 1894.4 | 2188.7 | 1830 | 2019.1 | 2092.2163 | 1968 | 2065 | 1840 | 2042.8066
| 1 | 5 | 2280 | 2160 | 1050 | 1190 | 953 | 2032.9013 | 1863.8 | 2179.9 | 1850 | 2078.9 | 2098.7302 | 1948 | 2065 | 1870 | 110
| 1 | 6 | 2330 | 2200 | 1050 | 1175 | 980 | 2007.5389 | 1863.7 | 2188.5 | 1870 | 2050.96 | 2161.2653 | 1997 | 2070 | 1920
| Mean |  | 2305 | 2168.3333 | 1056.6667 | 1158.3333 | 970 | 2046.5362 | 1886.6167 | 2199.3 | 1845 | 2039.005 | 2093.9143 | 1938.1667 | 2067.5 | 1856.6667
| Median |  | 2305 | 2160 | 1055 | 1160 | 963 | 2042.581 | 1892.5 | 2191.8 | 1845 | 2036.665 | 2093.1755 | 1951 | 2067.5 | 1850
| Std Dev. |  | 18.7083 | 27.1416 | 8.165 | 23.8048 | 22.2081 | 32.3278 | 19.3268 | 18.6191 | 18.7083 | 35.6645 | 38.1016 | 63.5655 | 2.7386 | 36.1478
| Rel.Std.Dev. |  | 0.8116% | 1.252% | 0.7727% | 2.055% | 2.289% | 1.58% | 1.024% | 0.8466% | 1.014% | 1.749% | 1.82% | 3.28% | 0.1325% | 1.947%
| PDM3 |  | 12.83% | 6.145% | -48.27% | -43.3% | -52.52% | 0.1826% | -7.646% | 7.661% | -9.683% | -0.1861% | 2.502% | -5.122% | 1.209% | -9.112%
| Z-Score (Absolute) |  | 1.37 | 0.75 | 4.24 | 3.78 | 4.67 | 0.2 | 0.51 | 0.89 | 0.7 | 0.17 | 0.42 | 0.17 | 0.3 | 0.65
| Analytical results for Ta in OREAS 464 (Certified Value 20.9 ppm) | [15 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 24.1 | 22.9 | 22.5 | 22.9 | 19.1134 | 20.4 | 15.1 | 11.2 | 16.57 | 20.5 |  |  |  |  | 1
| 1 | 2 | 23.9 | 23 | 22.5 | 22 | 18.8427 | 21 | 10.4 | 11.1 | 13.41 | 20.9 |  |  |  |  | 2
| 1 | 3 | 24.3 | 22.6 | 22.2 | 22.4 | 18.2164 | 20.8 | 12.1 | 11.6 | 16.65 | 19.6 |  |  |  |  | 16
| 1 | 4 | 23.7 | 22.3 | 22.8 | 22.9 | 18.8374 | 21.6 | 12.6 | 11.7 | 16.91 | 20.1 |  |  |  |  | 20.9121
| 1 | 5 | 23.1 | 23.3 | 23 | 23.3 | 18.1333 | 21 | 8.5 | 11.5 | 14.76 | 21.3 |  |  |  |  | 111
| 1 | 6 | 23.8 | 22.2 | 22.1 | 23.1 | 19.3882 | 20.2 | 14.2 | 11.5 | 13.75 | 20.9
| Mean |  | 23.8167 | 22.7167 | 22.5167 | 22.7667 | 18.7552 | 20.8333 | 12.15 | 11.4333 | 15.3417 | 20.55
| Median |  | 23.85 | 22.75 | 22.5 | 22.9 | 18.8401 | 20.9 | 12.35 | 11.5 | 15.665 | 20.7
| Std Dev. |  | 0.41191 | 0.42622 | 0.34303 | 0.48028 | 0.49405 | 0.49666 | 2.4288 | 0.23381 | 1.5674 | 0.61887
| Rel.Std.Dev. |  | 1.729% | 1.876% | 1.523% | 2.11% | 2.634% | 2.384% | 19.99% | 2.045% | 10.22% | 3.012%
| PDM3 |  | 13.89% | 8.629% | 7.673% | 8.868% | -10.31% | -0.3767% | -41.9% | -45.33% | -26.64% | -1.732%
| Z-Score (Absolute) |  | 1.03 | 0.67 | 0.6 | 0.68 | 0.64 | 0.05 | 2.81 | 3.05 | 1.76 | 0.05
| Analytical results for Tb in OREAS 464 (Certified Value 55 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 54 | 56.3 | 50.9 | 53 | 59.1 | 54.3175 | 45.2 | 56.08 | 33.6 | 57.66 | 84.8365 | 53.05 | 53.9 |  | 1
| 1 | 2 | 54.3 | 55.4 | 51.1 | 52.1 | 57.7 | 54.13 | 46.5 | 56.78 | 33 | 57.61 | 81.8399 | 53.5 | 54.7 |  | 3
| 1 | 3 | 54.2 | 50 | 50.4 | 52.3 | 58.1 | 53.2218 | 46 | 56.43 | 34.6 | 59.31 | 80.7973 | 54.8 | 55.6 |  | 16
| 1 | 4 | 54.4 | 55.2 | 50.3 | 52.2 | 59.1 | 53.8758 | 45.8 | 56 | 32.5 | 57.76 | 82.2319 | 52.95 | 55.1 |  | 54.8891
| 1 | 5 | 54.8 | 55.2 | 50.5 | 53.4 | 61.2 | 54.249 | 44.6 | 55.74 | 32.9 | 57.81 | 82.9475 | 53.5 | 55.2 |  | 112
| 1 | 6 | 53.4 | 56.3 | 49.8 | 52.2 | 59.4 | 52.8394 | 47.8 | 57.43 | 33.5 | 58.17 | 81.7471 | 55.45 | 54.2
| Mean |  | 54.1833 | 54.7333 | 50.5 | 52.5333 | 59.1 | 53.7722 | 45.9833 | 56.41 | 33.35 | 58.0533 | 82.4001 | 53.875 | 54.7833
| Median |  | 54.25 | 55.3 | 50.45 | 52.25 | 59.1 | 54.0029 | 45.9 | 56.255 | 33.25 | 57.785 | 82.0359 | 53.5 | 54.9
| Std Dev. |  | 0.46655 | 2.3746 | 0.46043 | 0.53541 | 1.2215 | 0.60606 | 1.1071 | 0.61703 | 0.73417 | 0.64643 | 1.3837 | 1.0153 | 0.64317
| Rel.Std.Dev. |  | 0.8611% | 4.338% | 0.9118% | 1.019% | 2.067% | 1.127% | 2.408% | 1.094% | 2.201% | 1.114% | 1.679% | 1.884% | 1.174%
| PDM3 |  | -1.286% | -0.2837% | -7.996% | -4.292% | 7.672% | -2.035% | -16.22% | 2.771% | -39.24% | 5.765% | 50.12% | -1.847% | -0.1926%
| Z-Score (Absolute) |  | 0 | 0.45 | 1.12 | 0.5 | 1.49 | 0.12 | 2.48 | 0.67 | 6.31 | 1.17 | 8.54 | 0.09 | 0.18
| Analytical results for Te in OREAS 464 (Certified Value 0.36 ppm) | [15 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.43 | 0.31 | 0.49 | 0.34 | 0.34 | 0.27334 | < 1 | < 0.5 | 0.48 | 0.3 | 0.2 |  |  |  | 1
| 1 | 2 | 0.42 | 0.33 | 0.58 | 0.32 | 0.32 | 0.28112 | < 1 | < 0.5 | 0.46 | 0.4 | 0.2 |  |  |  | 16
| 1 | 3 | 0.4 | 0.34 | 0.58 | 0.4 | 0.33 | 0.24571 | < 1 | < 0.5 | 0.46 | 0.4 | 0.2 |  |  |  | 16
| 1 | 4 | 0.41 | 0.32 | 0.54 | 0.37 | 0.33 | 0.23765 | < 1 | < 0.5 | 0.46 | 0.3 | 0.2 |  |  |  | 0.3625
| 1 | 5 | 0.42 | 0.38 | 0.56 | 0.41 | 0.33 | 0.23405 | < 1 | < 0.5 | 0.49 | 0.3 | 0.2 |  |  |  | 113
| 1 | 6 | 0.43 | 0.36 | 0.47 | 0.41 | 0.34 | 0.273 | < 1 | < 0.5 | 0.44 | 0.4 | 0.2
| Mean |  | 0.41833 | 0.34 | 0.53667 | 0.375 | 0.33167 | 0.25748 |  |  | 0.465 | 0.35 | 0.2
| Median |  | 0.42 | 0.335 | 0.55 | 0.385 | 0.33 | 0.25936 |  |  | 0.46 | 0.35 | 0.2
| Std Dev. |  | 0.01169 | 0.026077 | 0.046762 | 0.038341 | 0.0075277 | 0.020648 |  |  | 0.017607 | 0.054772 | 3.0405e-17
| Rel.Std.Dev. |  | 2.795% | 7.67% | 8.713% | 10.22% | 2.27% | 8.019% |  |  | 3.786% | 15.65% | 1.52e-14%
| PDM3 |  | 15.4% | -6.206% | 48.05% | 3.449% | -8.505% | -28.97% |  |  | 28.28% | -3.447% | -44.83%
| Z-Score (Absolute) |  | 0.67 | 0.1 | 1.84 | 0.25 | 0.18 | 0.91 | NA | NA | 1.13 | 0 | 1.48
| Analytical results for Th in OREAS 464 (Certified Value 506 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 870 | 390 | 556 | 383 | 525 | 511.13 | 544.1 | 81.5 | 524.97 | 563.3469 | 517 |  |  |  | 1
| 1 | 2 | 840 | 405 | 534 | 386 | 508 | 508.5451 | 547 | 80.4 | 530.88 | 561.0541 | 524 |  |  |  | 0
| 1 | 3 | 870 | 410 | 542 | 405 | 517 | 498.5196 | 564.4 | 85.4 | 549.57 | 562.7466 | 537 |  |  |  | 16
| 1 | 4 | 880 | 390 | 540 | 406 | 529 | 508.1129 | 551.6 | 83.6 | 522.57 | 563.8497 | 536 |  |  |  | 506.3839
| 1 | 5 | 850 | 419 | 563 | 413 | 543 | 512.1607 | 536.7 | 82.9 | 536.57 | 564.5312 | 531 |  |  |  | 114
| 1 | 6 | 860 | 405 | 546 | 402 | 528 | 498.5466 | 552.5 | 87.3 | 550.8 | 563.5296 | 527
| Mean |  | 861.6667 | 403.1667 | 546.8333 | 399.1667 | 525 | 506.1692 | 549.3833 | 83.5167 | 535.8933 | 563.1763 | 528.6667
| Median |  | 865 | 405 | 544 | 403.5 | 526.5 | 508.329 | 549.3 | 83.25 | 533.725 | 563.4383 | 529
| Std Dev. |  | 14.7196 | 11.4091 | 10.7781 | 11.9569 | 11.8491 | 6.1082 | 9.323 | 2.5309 | 12.097 | 1.1943 | 7.607
| Rel.Std.Dev. |  | 1.708% | 2.83% | 1.971% | 2.995% | 2.257% | 1.207% | 1.697% | 3.03% | 2.257% | 0.2121% | 1.439%
| PDM3 |  | 70.16% | -20.38% | 7.988% | -21.17% | 3.676% | -0.04242% | 8.491% | -83.51% | 5.827% | 11.22% | 4.4%
| Z-Score (Absolute) |  | 9.98 | 3.76 | 0.54 | 3.88 | 0.11 | 0.67 | 0.62 | 13.34 | 0.22 | 1.03 | 0
| Analytical results for Ti in OREAS 464 (Certified Value < 1 wt.%) | [15 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.899 | 0.89 | 0.462 | 0.412 | 0.383 | 0.71148 | 0.67 | 0.7085 | 0.656 | 0.891 | 0.55 | 1.09 |  |  | 1
| 1 | 2 | 0.863 | 0.87 | 0.448 | 0.402 | 0.367 | 0.71117 | 0.68 | 0.6161 | 0.674 | 0.8842 | 0.55 | 1.1 |  |  | 4
| 1 | 3 | 0.895 | 0.9 | 0.453 | 0.397 | 0.392 | 0.7496 | 0.68 | 0.5938 | 0.667 | 0.9083 | 0.52 | 1.14 |  |  | 16
| 1 | 4 | 0.871 | 0.89 | 0.426 | 0.415 | 0.401 | 0.74343 | 0.69 | 0.6138 | 0.681 | 0.9009 | 0.52 | 1.2 |  |  | < 1
| 1 | 5 | 0.837 | 0.93 | 0.446 | 0.42 | 0.369 | 0.6888 | 0.68 | 0.5722 | 0.685 | 0.856 | 0.53 | 1.17 |  |  | 115
| 1 | 6 | 0.866 | 0.92 | 0.428 | 0.422 | 0.401 | 0.74753 | 0.66 | 0.7105 | 0.679 | 0.8493 | 0.54 | 1.14
| Mean |  | 0.87183 | 0.9 | 0.44383 | 0.41133 | 0.3855 | 0.72533 | 0.67667 | 0.63582 | 0.67367 | 0.88162 | 0.535 | 1.14
| Median |  | 0.8685 | 0.895 | 0.447 | 0.4135 | 0.3875 | 0.72746 | 0.68 | 0.61495 | 0.6765 | 0.8876 | 0.535 | 1.14
| Std Dev. |  | 0.022807 | 0.021909 | 0.014176 | 0.0099532 | 0.015123 | 0.025045 | 0.010328 | 0.059239 | 0.010652 | 0.023996 | 0.013784 | 0.041473
| Rel.Std.Dev. |  | 2.616% | 2.434% | 3.194% | 2.42% | 3.923% | 3.453% | 1.526% | 9.317% | 1.581% | 2.722% | 2.576% | 3.638%
PDM3
| Z-Score (Absolute) |  | 0.66 | 0.75 | 0.77 | 0.88 | 0.97 | 0.17 | 0.01 | 0.13 | 0.01 | 0.69 | 0.47 | 1.56
| Analytical results for Tl in OREAS 464 (Certified Value 0.15 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 0.15 | 0.09 | 0.15 | 0.17 | 0.15 | 0.1429 | 0.14 | < 0.2 | 0.14 | 0.13 | 0.16333 | 0.15 | 0.1 |  | 1
| 1 | 2 | 0.15 | 0.1 | 0.16 | 0.15 | 0.14 | 0.1357 | 0.14 | 0.2 | 0.14 | 0.12 | 0.21759 | 0.15 | 0.1 |  | 17
| 1 | 3 | 0.15 | 0.05 | 0.15 | 0.17 | 0.14 | 0.15358 | 0.15 | < 0.2 | 0.15 | 0.11 | 0.21743 | 0.15 | 0.1 |  | 16
| 1 | 4 | 0.15 | 0.09 | 0.14 | 0.16 | 0.15 | 0.12752 | 0.15 | < 0.2 | 0.15 | 0.13 | 0.2291 | 0.15 | 0.1 |  | 0.14535
| 1 | 5 | 0.16 | 0.09 | 0.14 | 0.17 | 0.15 | 0.13365 | 0.15 | < 0.2 | 0.14 | 0.12 | 0.19834 | 0.15 | 0.1 |  | 116
| 1 | 6 | 0.16 | 0.08 | 0.13 | 0.15 | 0.15 | 0.15551 | 0.15 | < 0.2 | 0.14 | 0.11 | 0.21758 | 0.15 | 0.1
| Mean |  | 0.15333 | 0.083333 | 0.145 | 0.16167 | 0.14667 | 0.14147 | 0.14667 | 0.2 | 0.14333 | 0.12 | 0.20723 | 0.15 | 0.1
| Median |  | 0.15 | 0.09 | 0.145 | 0.165 | 0.15 | 0.1393 | 0.15 | 0.2 | 0.14 | 0.12 | 0.21751 | 0.15 | 0.1
| Std Dev. |  | 0.005164 | 0.017512 | 0.010488 | 0.0098319 | 0.005164 | 0.011269 | 0.005164 |  | 0.005164 | 0.0089443 | 0.023676 | 0 | 1.5202e-17
| Rel.Std.Dev. |  | 3.368% | 21.01% | 7.233% | 6.082% | 3.521% | 7.965% | 3.521% |  | 3.603% | 7.454% | 11.43% | 0% | 1.52e-14%
| PDM3 |  | 5.493% | -42.67% | -0.2402% | 11.23% | 0.9065% | -2.665% | 0.9065% | 37.6% | -1.387% | -17.44% | 42.57% | 3.2% | -31.2%
| Z-Score (Absolute) |  | 0.67 | 5.73 | 0.17 | 1.52 | 0 | 0.53 | 0 | 5.39 | 0.34 | 2.7 | 7.01 | 0.34 | 4.72
| Analytical results for Tm in OREAS 464 (Certified Value 3.3 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 3.84 | 3.63 | 3.27 | 3.3 | 3.52 | 3.087 | 2.6545 | 3.7 | 1.87 | 2.97 | 4.3939 | 3.4 | 3.02 |  | 1
| 1 | 2 | 3.75 | 3.61 | 3.25 | 3.29 | 3.46 | 3.2078 | 2.704 | 3.7 | 1.85 | 2.96 | 4.2954 | 3.35 | 3.04 |  | 5
| 1 | 3 | 3.77 | 3.21 | 3.21 | 3.26 | 3.45 | 3.0746 | 2.7742 | 3.8 | 1.97 | 2.92 | 4.1532 | 3.4 | 3.08 |  | 16
| 1 | 4 | 3.79 | 3.44 | 3.2 | 3.25 | 3.54 | 3.0566 | 2.6781 | 3.5 | 1.78 | 2.92 | 4.2428 | 3.35 | 3.16 |  | 3.2956
| 1 | 5 | 3.76 | 3.27 | 3.3 | 3.39 | 3.66 | 3.261 | 2.6198 | 3.6 | 1.82 | 3.09 | 4.2235 | 3.4 | 3.16 |  | 117
| 1 | 6 | 3.83 | 3.61 | 3.16 | 3.26 | 3.52 | 3.13 | 2.6725 | 3.7 | 1.75 | 2.88 | 4.3145 | 3.45 | 3.16
| Mean |  | 3.79 | 3.4617 | 3.2317 | 3.2917 | 3.525 | 3.1362 | 2.6839 | 3.6667 | 1.84 | 2.9567 | 4.2706 | 3.3917 | 3.1033
| Median |  | 3.78 | 3.525 | 3.23 | 3.275 | 3.52 | 3.1085 | 2.6753 | 3.7 | 1.835 | 2.94 | 4.2691 | 3.4 | 3.12
| Std Dev. |  | 0.037417 | 0.18595 | 0.051153 | 0.051929 | 0.075299 | 0.0816 | 0.052328 | 0.10328 | 0.07746 | 0.072847 | 0.083074 | 0.037639 | 0.065013
| Rel.Std.Dev. |  | 0.9872% | 5.372% | 1.583% | 1.578% | 2.136% | 2.602% | 1.95% | 2.817% | 4.21% | 2.464% | 1.945% | 1.11% | 2.095%
| PDM3 |  | 15% | 5.038% | -1.941% | -0.1205% | 6.96% | -4.839% | -18.56% | 11.26% | -44.17% | -10.29% | 29.58% | 2.914% | -5.835%
| Z-Score (Absolute) |  | 1.38 | 0.51 | 0.11 | 0 | 0.67 | 0.36 | 1.57 | 1.14 | 3.82 | 0.84 | 2.66 | 0.32 | 0.45
| Analytical results for U in OREAS 464 (Certified Value 17.6 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 17.2 | 17.9 | 17.8 | 18.4 | 19.6 | 16.6327 | 18.13 | 18.1 | 17.9 | 17.11 | 22.176 | 17.3 | 16.2 |  | 1
| 1 | 2 | 16.8 | 18.5 | 17.2 | 18.1 | 19.3 | 16.4923 | 18.13 | 18.2 | 18.1 | 17.4 | 21.2595 | 17.2 | 15.2 |  | 6
| 1 | 3 | 17.1 | 15.6 | 17.6 | 18 | 19.2 | 16.1846 | 18.11 | 18.1 | 18.6 | 17.44 | 21.6651 | 18 | 16.5 |  | 16
| 1 | 4 | 17.4 | 17.1 | 17.1 | 18 | 19.6 | 16.2618 | 18.76 | 18 | 17.9 | 17.15 | 21.6444 | 17.4 | 16.4 |  | 17.6356
| 1 | 5 | 17.1 | 18.1 | 18.3 | 18.7 | 20 | 16.5144 | 18.33 | 17.8 | 18.5 | 17.94 | 21.6024 | 17.1 | 15.7 |  | 118
| 1 | 6 | 17.6 | 17.6 | 17.4 | 18.1 | 19.4 | 16.4574 | 17.52 | 18.1 | 18.1 | 17.31 | 22.0646 | 17.5 | 15.6
| Mean |  | 17.2 | 17.4667 | 17.5667 | 18.2167 | 19.5167 | 16.4239 | 18.1633 | 18.05 | 18.1833 | 17.3917 | 21.7353 | 17.4167 | 15.9333
| Median |  | 17.15 | 17.75 | 17.5 | 18.1 | 19.5 | 16.4748 | 18.13 | 18.1 | 18.1 | 17.355 | 21.6548 | 17.35 | 15.95
| Std Dev. |  | 0.27568 | 1.0289 | 0.44121 | 0.27869 | 0.28577 | 0.16801 | 0.40068 | 0.13784 | 0.29944 | 0.29916 | 0.33461 | 0.31885 | 0.51251
| Rel.Std.Dev. |  | 1.603% | 5.891% | 2.512% | 1.53% | 1.464% | 1.023% | 2.206% | 0.7637% | 1.647% | 1.72% | 1.539% | 1.831% | 3.217%
| PDM3 |  | -2.47% | -0.9577% | -0.3906% | 3.295% | 10.67% | -6.871% | 2.993% | 2.35% | 3.106% | -1.383% | 23.25% | -1.241% | -9.652%
| Z-Score (Absolute) |  | 1.02 | 0 | 0.44 | 0.45 | 2.67 | 2.26 | 0.52 | 0.33 | 0.55 | 0.71 | 6.2 | 0.67 | 3.04
| Analytical results for V in OREAS 464 (Certified Value 207 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 199 | 187 | 111 | 117 | 106 | 204.9886 | 217.6 | 229 | 210 | 253 | 205.715 | 142 | 220 |  | 1
| 1 | 2 | 201 | 191 | 110 | 117 | 108 | 205.7898 | 223 | 220 | 209 | 258 | 207.297 | 149 | 220 |  | 7
| 1 | 3 | 198 | 190 | 110 | 117 | 107 | 201.1238 | 223.2 | 232 | 213 | 248 | 207.06 | 139 | 225 |  | 16
| 1 | 4 | 203 | 194 | 110 | 118 | 112 | 204.4475 | 219.5 | 228 | 202 | 243 | 208.287 | 147 | 230 |  | 207.0274
| 1 | 5 | 195 | 190 | 109 | 122 | 107 | 202.3107 | 217.5 | 226 | 205 | 239 | 207.559 | 145 | 225 |  | 119
| 1 | 6 | 204 | 206 | 111 | 119 | 109 | 199.965 | 215 | 225 | 206 | 242 | 207.898 | 151 | 220
| Mean |  | 200 | 193 | 110.1667 | 118.3333 | 108.1667 | 203.1042 | 219.3 | 226.6667 | 207.5 | 247.1667 | 207.3027 | 145.5 | 223.3333
| Median |  | 200 | 190.5 | 110 | 117.5 | 107.5 | 203.3791 | 218.55 | 227 | 207.5 | 245.5 | 207.428 | 146 | 222.5
| Std Dev. |  | 3.3466 | 6.7528 | 0.75277 | 1.9664 | 2.137 | 2.3227 | 3.2729 | 4.0825 | 3.937 | 7.2503 | 0.89091 | 4.4609 | 4.0825
| Rel.Std.Dev. |  | 1.673% | 3.499% | 0.6833% | 1.662% | 1.976% | 1.144% | 1.492% | 1.801% | 1.897% | 2.933% | 0.4298% | 3.066% | 1.828%
| PDM3 |  | -3.394% | -6.776% | -46.79% | -42.84% | -47.75% | -1.895% | 5.928% | 9.486% | 0.2283% | 19.39% | 0.133% | -29.72% | 7.876%
| Z-Score (Absolute) |  | 0.1 | 0.42 | 3.1 | 2.85 | 3.16 | 0 | 0.54 | 0.79 | 0.15 | 1.47 | 0.14 | 1.92 | 0.67
| Analytical results for W in OREAS 464 (Certified Value 2.37 ppm) | [15 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.9 | 2.5 | 3.2 | 2.7 | 2.7 | 2.2315 | 2 | 2 | 2.1 | 1.8 | 2 |  |  |  | 1
| 1 | 2 | 2.8 | 2.5 | 3.2 | 2.7 | 2.6 | 2.2597 | 2 | 2 | 1.9 | 1.9 | 2 |  |  |  | 8
| 1 | 3 | 2.8 | 2.3 | 3.1 | 2.5 | 2.7 | 2.0179 | 2 | 2 | 2.1 | 2.1 | 2.5 |  |  |  | 16
| 1 | 4 | 2.8 | 2.6 | 3 | 2.6 | 2.7 | 1.9838 | 2 | 2 | 2 | 1.9 | 2.5 |  |  |  | 2.3747
| 1 | 5 | 2.8 | 2.2 | 3.2 | 2.8 | 2.7 | 2.0233 | 2 | 2 | 2.1 | 1.8 | 2.5 |  |  |  | 120
| 1 | 6 | 2.8 | 2.5 | 3 | 2.7 | 2.7 | 2.3131 | 2 | 2 | 2.1 | 1.8 | 2.5
| Mean |  | 2.8167 | 2.4333 | 3.1167 | 2.6667 | 2.6833 | 2.1382 | 2 | 2 | 2.05 | 1.8833 | 2.3333
| Median |  | 2.8 | 2.5 | 3.15 | 2.7 | 2.7 | 2.1274 | 2 | 2 | 2.1 | 1.85 | 2.5
| Std Dev. |  | 0.040825 | 0.15055 | 0.098319 | 0.10328 | 0.040825 | 0.14534 | 0 | 0 | 0.083666 | 0.1169 | 0.2582
| Rel.Std.Dev. |  | 1.449% | 6.187% | 3.155% | 3.873% | 1.521% | 6.797% | 0% | 0% | 4.081% | 6.207% | 11.07%
| PDM3 |  | 18.61% | 2.47% | 31.25% | 12.3% | 13% | -9.959% | -15.78% | -15.78% | -13.67% | -20.69% | -1.741%
| Z-Score (Absolute) |  | 0.98 | 0.2 | 1.58 | 0.67 | 0.71 | 0.39 | 0.67 | 0.67 | 0.57 | 0.91 | 0
| Analytical results for Y in OREAS 464 (Certified Value 449 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 421 | 403 | 470 | 222 | 467 | 434.2993 | 262 | 470.1 | 157 | 449.04 | 516.4307 | 418.7 | 425 |  | 1
| 1 | 2 | 421 | 395 | 477 | 220 | 464 | 451.2891 | 261 | 464.2 | 152 | 449.43 | 516.4215 | 419.2 | 436 |  | 0
| 1 | 3 | 420 | 380 | 462 | 221 | 456 | 449.883 | 282 | 471.7 | 166 | 463.29 | 513.8517 | 435.5 | 442 |  | 16
| 1 | 4 | 420 | 385 | 467 | 226 | 471 | 430.9344 | 267 | 461.9 | 157 | 453.05 | 517.5494 | 416.8 | 435 |  | 448.8679
| 1 | 5 | 421 | 391 | 474 | 230 | 489 | 427.3796 | 252 | 467.5 | 155 | 468.73 | 518.6566 | 424.7 | 439 |  | 121
| 1 | 6 | 421 | 412 | 454 | 225 | 469 | 429.2373 | 282 | 463.6 | 166 | 455.7 | 516.0025 | 432 | 438
| Mean |  | 420.6667 | 394.3333 | 467.3333 | 224 | 469.3333 | 437.1705 | 267.6667 | 466.5 | 158.8333 | 456.54 | 516.4854 | 424.4833 | 435.8333
| Median |  | 421 | 393 | 468.5 | 223.5 | 468 | 432.6169 | 264.5 | 465.85 | 157 | 454.375 | 516.4261 | 421.95 | 437
| Std Dev. |  | 0.5164 | 11.7587 | 8.3825 | 3.7417 | 10.9666 | 10.6477 | 12.1106 | 3.8951 | 5.8452 | 7.9227 | 1.6139 | 7.7241 | 5.8452
| Rel.Std.Dev. |  | 0.1228% | 2.982% | 1.794% | 1.67% | 2.337% | 2.436% | 4.525% | 0.835% | 3.68% | 1.735% | 0.3125% | 1.82% | 1.341%
| PDM3 |  | -6.283% | -12.15% | 4.114% | -50.1% | 4.559% | -2.606% | -40.37% | 3.928% | -64.61% | 1.709% | 15.06% | -5.432% | -2.904%
| Z-Score (Absolute) |  | 0.32 | 0.89 | 0.67 | 4.53 | 0.72 | 0.03 | 3.6 | 0.66 | 5.93 | 0.44 | 1.73 | 0.24 | 0
| Analytical results for Yb in OREAS 464 (Certified Value 14.3 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 14.95 | 14.6 | 13.55 | 13.8 | 14.1 | 13.5134 | 10.6 | 15.2 | 10.2 | 13.21 | 19.5926 | 15.5 | 14.7 |  | 1
| 1 | 2 | 14.45 | 14.4 | 13.75 | 13.5 | 14 | 13.18 | 11.3377 | 15.7 | 10.1 | 13.21 | 18.3324 | 15.65 | 14.8 |  | 9
| 1 | 3 | 14.65 | 12.75 | 13.3 | 13.5 | 13.85 | 12.638 | 11.3059 | 15.2 | 10.8 | 13.49 | 18.259 | 16.2 | 15 |  | 16
| 1 | 4 | 14.4 | 14.1 | 13.15 | 13.45 | 14.3 | 12.9755 | 11.0977 | 15.7 | 10.7 | 13.11 | 18.4202 | 15.65 | 15.2 |  | 14.277
| 1 | 5 | 14.35 | 14.2 | 13.85 | 13.85 | 14.7 | 13.3024 | 10.5895 | 15.5 | 10 | 14 | 18.5084 | 16.05 | 14.8 |  | 122
| 1 | 6 | 14.8 | 14.55 | 13.65 | 13.75 | 14.1 | 12.8498 | 10.9526 | 15.7 | 10.7 | 13.12 | 19.6124 | 16 | 14.9
| Mean |  | 14.6 | 14.1 | 13.5417 | 13.6417 | 14.175 | 13.0765 | 10.9806 | 15.5 | 10.4167 | 13.3567 | 18.7875 | 15.8417 | 14.9
| Median |  | 14.55 | 14.3 | 13.6 | 13.625 | 14.1 | 13.0777 | 11.0252 | 15.6 | 10.45 | 13.21 | 18.4643 | 15.825 | 14.85
| Std Dev. |  | 0.24083 | 0.6892 | 0.2691 | 0.17725 | 0.29623 | 0.31829 | 0.33037 | 0.24495 | 0.35449 | 0.34396 | 0.63685 | 0.27824 | 0.17889
| Rel.Std.Dev. |  | 1.65% | 4.888% | 1.987% | 1.299% | 2.09% | 2.434% | 3.009% | 1.58% | 3.403% | 2.575% | 3.39% | 1.756% | 1.201%
| PDM3 |  | 2.263% | -1.239% | -5.15% | -4.45% | -0.7141% | -8.408% | -23.09% | 8.567% | -27.04% | -6.446% | 31.59% | 10.96% | 4.364%
| Z-Score (Absolute) |  | 0.42 | 0.24 | 0.42 | 0.34 | 0 | 0.8 | 2.47 | 1.15 | 2.93 | 0.67 | 3.78 | 1.42 | 0.66
| Analytical results for Zn in OREAS 464 (Certified Value 1128 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 1080 | 1220 | 516 | 528 | 467 | 1160.7229 | 1087.2889 | 1173 | 1110 | 1094 | 1204.47 | 1010 | 1115 | 1090 | 1
| 1 | 2 | 1100 | 1210 | 504 | 520 | 469 | 1171.2183 | 1135.4791 | 1155 | 1110 | 1080 | 1197.1692 | 1108 | 1110 | 1080 | 10
| 1 | 3 | 1080 | 1230 | 514 | 526 | 474 | 1180.2368 | 1109.2882 | 1170 | 1100 | 1062 | 1192.7196 | 1058 | 1115 | 1090 | 16
| 1 | 4 | 1090 | 1280 | 506 | 536 | 472 | 1175.7576 | 1107.2715 | 1174 | 1110 | 1080 | 1204.686 | 1068 | 1110 | 1090 | 1128.0782
| 1 | 5 | 1050 | 1200 | 504 | 548 | 471 | 1164.2548 | 1097.9691 | 1152 | 1100 | 1069 | 1194.8688 | 1022 | 1115 | 1100 | 123
| 1 | 6 | 1080 | 1280 | 507 | 540 | 488 | 1176.2111 | 1079.9697 | 1214 | 1110 | 1101 | 1278.0288 | 1099 | 1110 | 1110
| Mean |  | 1080 | 1236.6667 | 508.5 | 533 | 473.5 | 1171.4002 | 1102.8778 | 1173 | 1106.6667 | 1081 | 1211.9904 | 1060.8333 | 1112.5 | 1093.3333
| Median |  | 1080 | 1225 | 506.5 | 532 | 471.5 | 1173.488 | 1102.6203 | 1171.5 | 1110 | 1080 | 1200.8196 | 1063 | 1112.5 | 1090
| Std Dev. |  | 16.7332 | 35.0238 | 5.2058 | 10.2567 | 7.5033 | 7.554 | 19.5711 | 22.163 | 5.164 | 14.6697 | 32.7266 | 39.5799 | 2.7386 | 10.328
| Rel.Std.Dev. |  | 1.549% | 2.832% | 1.024% | 1.924% | 1.585% | 0.6449% | 1.775% | 1.889% | 0.4666% | 1.357% | 2.7% | 3.731% | 0.2462% | 0.9446%
| PDM3 |  | -4.262% | 9.626% | -54.92% | -52.75% | -58.03% | 3.84% | -2.234% | 3.982% | -1.898% | -4.173% | 7.439% | -5.961% | -1.381% | -3.08%
| Z-Score (Absolute) |  | 0.23 | 1.8 | 7.65 | 7.33 | 8.14 | 0.95 | 0.06 | 0.87 | 0.11 | 0.22 | 1.31 | 0.48 | 0.19 | 0.06
| Analytical results for Zr in OREAS 464 (Certified Value 210 ppm) | [15 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 243 | 221 | 250 | 244 | 237 | 187.7672 | 216 | 104 | 186.7 | 122 | 110 | 407 |  |  | 1
| 1 | 2 | 238 | 222 | 244 | 240 | 236 | 189.4418 | 202 | 108 | 170.8 | 118 | 95 | 412 |  |  | 11
| 1 | 3 | 241 | 210 | 234 | 240 | 231 | 187.5 | 207 | 103 | 186.8 | 119 | 95 | 406 |  |  | 16
| 1 | 4 | 240 | 222 | 234 | 246 | 240 | 188.4863 | 204 | 106 | 192.8 | 115 | 160 | 406 |  |  | 209.7096
| 1 | 5 | 234 | 234 | 243 | 251 | 246 | 186.582 | 196 | 107 | 188.8 | 106 | 100 | 409 |  |  | 124
| 1 | 6 | 238 | 233 | 231 | 242 | 238 | 189.3831 | 218 | 109 | 181.7 | 117 | 165 | 419
| Mean |  | 239 | 223.6667 | 239.3333 | 243.8333 | 238 | 188.1934 | 207.1667 | 106.1667 | 184.6 | 116.1667 | 120.8333 | 409.8333
| Median |  | 239 | 222 | 238.5 | 243 | 237.5 | 188.1267 | 205.5 | 106.5 | 186.75 | 117.5 | 105 | 408
| Std Dev. |  | 3.0984 | 8.8694 | 7.4207 | 4.2151 | 4.9396 | 1.1242 | 8.4479 | 2.3166 | 7.6548 | 5.4924 | 32.7745 | 5.0365
| Rel.Std.Dev. |  | 1.296% | 3.965% | 3.101% | 1.729% | 2.075% | 0.5973% | 4.078% | 2.182% | 4.147% | 4.728% | 27.12% | 1.229%
| PDM3 |  | 13.97% | 6.655% | 14.13% | 16.27% | 13.49% | -10.26% | -1.213% | -49.37% | -11.97% | -44.61% | -42.38% | 95.43%
| Z-Score (Absolute) |  | 0.58 | 0.2 | 0.58 | 0.69 | 0.55 | 0.66 | 0.2 | 2.67 | 0.68 | 2.37 | 2.31 | 4.74
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.