## Abbreviations
Table 1. Abbreviations used for OREAS 465
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*OES/MS | lithium borate fusion with ICP-OES or ICP-MS finish as appropriate
BF*XRF | lithium borate fusion with XRF finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 465
Laboratory Name and Location
ALS, Brisbane, QLD, Australia
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Perth, WA, Australia
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
Nagrom, Perth, WA, Australia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
SGS Vostok Limited, Chita, Russian Federation
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
UIS Analytical Services, Centurion , South Africa
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 465
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Borate Fusion XRF
Ce, Cerium (wt.%) | 3.9692 | 0.036778 | 3.9371 | 4.0013 | 3.9411 | 3.9972
Fe, Iron (wt.%) | 34.9442 | 0.91532 | 33.7283 | 36.16 | 34.7098 | 35.1785
La, Lanthanum (wt.%) | 2.4205 | 0.028158 | 2.3969 | 2.4441 | 2.4036 | 2.4374
Nd, Neodymium (wt.%) | 1.1807 | 0.087725 | 1.0949 | 1.2665 | 1.1589 | 1.2026
Pr, Praseodymium (ppm) | 3752.3079 | 119.0932 | 3610.2535 | 3894.3622 | 3602.5822 | 3902.0335
Sm, Samarium (ppm) | 1414.1286 | 116.9916 | 1293.1805 | 1535.0767 | IND | IND
Th, Thorium (ppm) | 792.0682 | 112.6968 | 665.3729 | 918.7635 | IND | IND
Y, Yttrium (ppm) | 502.0015 | 35.4539 | 483.7335 | 520.2694 | IND | IND
Thermogravimetry
LOI, Loss On Ignition @1000°C (wt.%) | 0.8238 | 0.13277 | 0.6973 | 0.9503 | 0.77398 | 0.87362
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 6.5976 | 0.26121 | 6.395 | 6.8003 | 6.4744 | 6.7209
Ba, Barium (ppm) | 4397.0592 | 464.1253 | 4126.9659 | 4667.1526 | 4277.8738 | 4516.2446
Be, Beryllium (ppm) | 13.246 | 0.89319 | 12.5329 | 13.9592 | IND | IND
Bi, Bismuth (ppm) | 16.9218 | 1.0416 | 15.9693 | 17.8743 | 16.0937 | 17.7499
Ca, Calcium (wt.%) | 0.89974 | 0.059978 | 0.86134 | 0.93813 | 0.8706 | 0.92888
Ce, Cerium (wt.%) | 3.9525 | 0.13505 | 3.87 | 4.035 | 3.8782 | 4.0268
Cr, Chromium (ppm) | 544.1822 | 41.9089 | 516.2631 | 572.1013 | 530.4481 | 557.9163
Cs, Cesium (ppm) | < 0.1 | IND | IND | IND | IND | IND
Dy, Dysprosium (ppm) | 216.7238 | 12.5779 | 209.3619 | 224.0858 | 210.3614 | 223.0863
Er, Erbium (ppm) | 50.459 | 3.1388 | 48.6339 | 52.284 | 48.591 | 52.3269
Eu, Europium (ppm) | 285.6091 | 11.4732 | 279.6931 | 291.525 | 276.9218 | 294.2963
Fe, Iron (wt.%) | 34.7122 | 0.63354 | 34.162 | 35.2624 | 33.824 | 35.6005
Ga, Gallium (ppm) | 213.7519 | 42.8101 | 166.031 | 261.4728 | 207.9624 | 219.5414
Gd, Gadolinium (ppm) | 584.393 | 30.6922 | 567.5349 | 601.2512 | 569.1003 | 599.6858
Hf, Hafnium (ppm) | 41.4062 | 7.2079 | 35.8327 | 46.9798 | 38.5194 | 44.2931
Ho, Holmium (ppm) | 27.7013 | 2.1321 | 26.5363 | 28.8663 | 26.7465 | 28.6561
In, Indium (ppm) | 3.4687 | 0.22668 | 3.3157 | 3.6217 | 2.9353 | 4.0021
La, Lanthanum (wt.%) | 2.4128 | 0.082359 | 2.368 | 2.4577 | 2.3588 | 2.4669
Lu, Lutetium (ppm) | 2.1042 | 0.12164 | 2.0568 | 2.1515 | 1.9938 | 2.2145
Mg, Magnesium (wt.%) | 0.39174 | 0.020505 | 0.3772 | 0.40628 | 0.37961 | 0.40388
Mn, Manganese (wt.%) | 0.26252 | 0.022883 | 0.24651 | 0.27853 | 0.254 | 0.27105
Mo, Molybdenum (ppm) | 113.945 | 5.6664 | 108.9385 | 118.9515 | 108.1379 | 119.7522
Nb, Niobium (ppm) | 4680.1729 | 301.2312 | 4462.0762 | 4898.2695 | 4432.7541 | 4927.5916
Nd, Neodymium (wt.%) | 1.1764 | 0.050429 | 1.1479 | 1.2049 | 1.1518 | 1.2009
P, Phosphorus (wt.%) | 3.8122 | 0.12235 | 3.7044 | 3.92 | 3.6869 | 3.9375
Pb, Lead (ppm) | 506.1513 | 18.6321 | 496.053 | 516.2496 | 480.2575 | 532.0452
Pr, Praseodymium (ppm) | 3771.8124 | 164.023 | 3677.1027 | 3866.5222 | 3670.8846 | 3872.7403
S, Sulphur (ppm) | 1941 | 132.44 | 1787.0275 | 2094.9725 | IND | IND
Si, Silicon (wt.%) | 1.5341 | 0.04334 | 1.4849 | 1.5832 | 1.4922 | 1.5759
Sm, Samarium (ppm) | 1361.0146 | 36.1739 | 1344.9034 | 1377.1259 | 1316.6249 | 1405.4044
Sn, Tin (ppm) | 135.8085 | 20.2242 | 123.62 | 147.997 | 128.8209 | 142.7961
Sr, Strontium (ppm) | 5203.8067 | 182.0822 | 5085.291 | 5322.3224 | 5102.2144 | 5305.3989
Ta, Tantalum (ppm) | 78.6692 | 5.0443 | 74.8581 | 82.4802 | 74.1557 | 83.1827
Tb, Terbium (ppm) | 56.5714 | 3.1201 | 54.8819 | 58.261 | 54.2424 | 58.9005
Th, Thorium (ppm) | 865.5289 | 42.8179 | 840.4421 | 890.6157 | 844.963 | 886.0948
Ti, Titanium (wt.%) | 6.3017 | 0.18382 | 6.1628 | 6.4406 | 6.1433 | 6.46
Tm, Thulium (ppm) | 4.5221 | 0.26208 | 4.3966 | 4.6475 | 4.2911 | 4.753
U, Uranium (ppm) | 13.6006 | 0.40075 | 13.4105 | 13.7906 | 13.0474 | 14.1537
V, Vanadium (ppm) | 534.1117 | 35.5843 | 511.6763 | 556.5471 | 517.5574 | 550.666
W, Tungsten (ppm) | 7.52 | 1.2382 | 6.5067 | 8.5333 | IND | IND
Y, Yttrium (ppm) | 523.8165 | 33.1896 | 503.8552 | 543.7779 | 503.2625 | 544.3705
Yb, Ytterbium (ppm) | 18.9818 | 0.72455 | 18.6066 | 19.3569 | 17.908 | 20.0555
Zr, Zirconium (ppm) | 1879.6006 | 203.1451 | 1687.9891 | 2071.212 | 1763.5358 | 1995.6653
4-Acid Digestion
Ag, Silver (ppm) | 5.4832 | 0.35318 | 5.2957 | 5.6707 | 5.0725 | 5.8939
Al, Aluminium (wt.%) | 6.2092 | 0.39938 | 5.9969 | 6.4214 | 6.02 | 6.3983
Ba, Barium (ppm) | 4359.1998 | 157.9633 | 4270.3098 | 4448.0899 | 4291.3815 | 4427.0182
Be, Beryllium (ppm) | 11.6226 | 0.75763 | 11.1745 | 12.0708 | 11.1577 | 12.0876
Bi, Bismuth (ppm) | 17.2678 | 1.0295 | 16.6833 | 17.8523 | 16.7147 | 17.8209
Ca, Calcium (wt.%) | 0.87181 | 0.061686 | 0.82989 | 0.91373 | 0.8495 | 0.89412
Cd, Cadmium (ppm) | 1.203 | 0.14232 | 1.1247 | 1.2814 | 1.0846 | 1.3215
Ce, Cerium (wt.%) | 3.9056 | 0.14877 | 3.7644 | 4.0469 | 3.8132 | 3.998
Co, Cobalt (ppm) | 18.7017 | 1.2612 | 18.1006 | 19.3028 | 18.0695 | 19.3339
Cs, Cesium (ppm) | < 0.1 | IND | IND | IND | IND | IND
Cu, Copper (ppm) | 127.5667 | 5.3201 | 125.4927 | 129.6406 | 122.7764 | 132.357
Dy, Dysprosium (ppm) | 215.1915 | 9.2587 | 209.9341 | 220.4489 | 209.7212 | 220.6617
Er, Erbium (ppm) | 47.265 | 2.5083 | 45.7948 | 48.7352 | 45.535 | 48.995
Eu, Europium (ppm) | 282.2001 | 11.8075 | 274.5643 | 289.836 | 275.1506 | 289.2497
Fe, Iron (wt.%) | 29.5527 | 2.0248 | 28.3775 | 30.7279 | 28.9195 | 30.1858
Ga, Gallium (ppm) | 188.3139 | 30.7797 | 157.1232 | 219.5047 | 175.2718 | 201.3561
Gd, Gadolinium (ppm) | 581.0732 | 35.8098 | 555.7932 | 606.3531 | 569.7111 | 592.4353
Hf, Hafnium (ppm) | 14.4049 | 1.9332 | 13.7375 | 15.0723 | 13.0278 | 15.782
Ho, Holmium (ppm) | 26.7772 | 0.93535 | 26.335 | 27.2194 | 25.8621 | 27.6923
In, Indium (ppm) | 3.1831 | 0.20817 | 3.0596 | 3.3066 | 3.0492 | 3.3171
La, Lanthanum (wt.%) | 2.2739 | 0.092966 | 2.1545 | 2.3933 | 2.2188 | 2.329
Li, Lithium (ppm) | 3.0419 | 0.332 | 2.8029 | 3.2808 | 2.8595 | 3.2242
Lu, Lutetium (ppm) | 1.7206 | 0.28169 | 1.5379 | 1.9034 | 1.6245 | 1.8168
Mg, Magnesium (wt.%) | 0.37383 | 0.0261 | 0.35649 | 0.39117 | 0.36211 | 0.38555
Mn, Manganese (wt.%) | 0.19771 | 0.022452 | 0.18081 | 0.21461 | 0.1913 | 0.20412
Mo, Molybdenum (ppm) | 98.0101 | 10.1481 | 91.5029 | 104.5172 | 95.3702 | 100.65
Na, Sodium (wt.%) | < 0.2 | IND | IND | IND | IND | IND
Nd, Neodymium (wt.%) | 1.0951 | 0.043014 | 1.0529 | 1.1374 | 1.041 | 1.1493
Ni, Nickel (ppm) | 105.8075 | 13.5473 | 95.2186 | 116.3963 | 95.2108 | 116.4041
P, Phosphorus (wt.%) | 3.1508 | 0.37551 | 2.8297 | 3.4719 | 3.0613 | 3.2404
Pb, Lead (ppm) | 572.5402 | 16.0717 | 562.2869 | 582.7935 | 561.4835 | 583.5969
Pr, Praseodymium (ppm) | 3669.9075 | 192.7586 | 3487.8661 | 3851.9488 | 3561.4058 | 3778.4092
Rb, Rubidium (ppm) | 0.42577 | 0.054292 | 0.39679 | 0.45475 | 0.34719 | 0.50435
Re, Rhenium (ppm) | < 0.05 | IND | IND | IND | IND | IND
Sc, Scandium (ppm) | 148.5346 | 9.0796 | 143.1789 | 153.8904 | 143.4019 | 153.6674
Sm, Samarium (ppm) | 1306.679 | 80.8176 | 1230.3128 | 1383.0453 | 1240.8497 | 1372.5084
Sr, Strontium (wt.%) | 0.50533 | 0.02625 | 0.48743 | 0.52322 | 0.4902 | 0.52045
Tb, Terbium (ppm) | 57.0775 | 3.3289 | 55.0474 | 59.1075 | 55.404 | 58.751
Te, Tellurium (ppm) | < 1 | IND | IND | IND | IND | IND
Th, Thorium (ppm) | 804.6092 | 149.6429 | 693.3635 | 915.8549 | 772.6511 | 836.5673
Tl, Thallium (ppm) | 0.08732 | 0.0089241 | 0.081895 | 0.092746 | IND | IND
Tm, Thulium (ppm) | 3.8238 | 0.2742 | 3.6494 | 3.9982 | 3.7006 | 3.947
U, Uranium (ppm) | 12.6051 | 0.67854 | 12.2488 | 12.9615 | 12.131 | 13.0792
V, Vanadium (ppm) | 426.604 | 68.9831 | 380.402 | 472.8061 | 410.4007 | 442.8074
Y, Yttrium (ppm) | 477.799 | 32.3574 | 455.4187 | 500.1793 | 463.9448 | 491.6532
Yb, Ytterbium (ppm) | 14.9004 | 0.95238 | 14.2593 | 15.5415 | 14.2704 | 15.5304
Zn, Zinc (ppm) | 921.4187 | 133.0092 | 841.4058 | 1001.4317 | 893.8732 | 948.9642

## Indicative Values
Table 3. Indicative Values for OREAS 465
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Borate Fusion XRF
Al | wt.% | 6.6786 | Lu | ppm | < 20 | Sr | wt.% | 0.51018
Ba | ppm | 4784.2962 | Mg | wt.% | 0.40018 | Ta | ppm | < 100
Ca | wt.% | 0.93804 | Mn | wt.% | 0.28559 | Tb | ppm | 56.6797
Cr | ppm | 502.908 | Na | wt.% | 0.12751 | Ti | wt.% | 6.3505
Dy | ppm | 287.0675 | Nb | ppm | 4649.3621 | Tm | ppm | < 10
Er | ppm | 53.928 | Ni | ppm | 157.1709 | U | ppm | < 100
Eu | ppm | 302.2714 | P | wt.% | 3.8275 | V | ppm | 609.1754
Gd | ppm | 755.0574 | Pb | ppm | 603.4163 | W | ppm | < 100
Hf | ppm | < 100 | Si | wt.% | 1.4973 | Yb | ppm | 69.8249
Ho | ppm | 26.1894 | Sn | ppm | 170.6574 | Zn | ppm | 1205.0129
K | wt.% | 0.012279 | S | wt.% | 0.19422 | Zr | ppm | 1937.1237
Thermogravimetry
H2O- | wt.% | 0.31167
Borate / Peroxide Fusion ICP
Ag | ppm | 6.5 | K | wt.% | 0.099038 | Sc | ppm | 170.6667
As | ppm | 419.6712 | Li | ppm | 3.3333 | Se | ppm | 33.5833
B | ppm | 23.3333 | Na | wt.% | 0.077808 | Te | ppm | < 1
Cd | ppm | 1.0389 | Ni | ppm | 100.7898 | Tl | ppm | < 0.5
Co | ppm | 22.9817 | Rb | ppm | 0.98192 | Zn | ppm | 1097.0833
Cu | ppm | 128.3315 | Re | ppm | < 0.1
Ge | ppm | 60 | Sb | ppm | 0.98407
4-Acid Digestion
As | ppm | 31.7468 | S | ppm | 1103.0972 | Ti | wt.% | 1.3009
Cr | ppm | 356.9204 | Sb | ppm | 0.5746 | W | ppm | 2.397
Ge | ppm | 23.5885 | Se | ppm | 35.4167 | Zr | ppm | 367.3325
K | wt.% | 0.023204 | Sn | ppm | 48.7534
Nb | ppm | 398.9111 | Ta | ppm | 50.1017

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 465
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Borate Fusion XRF
Ce, wt.% | 3.9692 | 0.036778 | 3.8956 | 4.0427 | 3.8588 | 4.0795 | 0.9266% | 1.853% | 2.78% | 3.7707 | 4.1676
Fe, wt.% | 34.9442 | 0.91532 | 33.1135 | 36.7748 | 32.1982 | 37.6901 | 2.619% | 5.239% | 7.858% | 33.197 | 36.6914
La, wt.% | 2.4205 | 0.028158 | 2.3642 | 2.4768 | 2.3361 | 2.505 | 1.163% | 2.327% | 3.49% | 2.2995 | 2.5415
Nd, wt.% | 1.1807 | 0.087725 | 1.0053 | 1.3562 | 0.91757 | 1.4439 | 7.43% | 14.86% | 22.29% | 1.1217 | 1.2398
Pr, ppm | 3752.3079 | 119.0932 | 3514.1215 | 3990.4942 | 3395.0283 | 4109.5874 | 3.174% | 6.348% | 9.522% | 3564.6925 | 3939.9232
Sm, ppm | 1414.1286 | 116.9916 | 1180.1455 | 1648.1118 | 1063.1539 | 1765.1033 | 8.273% | 16.55% | 24.82% | 1343.4222 | 1484.8351
Th, ppm | 792.0682 | 112.6968 | 566.6745 | 1017.4618 | 453.9777 | 1130.1587 | 14.23% | 28.46% | 42.68% | 752.4648 | 831.6716
Y, ppm | 502.0015 | 35.4539 | 431.0937 | 572.9092 | 395.6398 | 608.3631 | 7.063% | 14.13% | 21.19% | 476.9014 | 527.1015
Thermogravimetry
LOI, wt.% | 0.8238 | 0.13277 | 0.55826 | 1.0893 | 0.4255 | 1.2221 | 16.12% | 32.23% | 48.35% | 0.78261 | 0.86499
Borate / Peroxide Fusion ICP
Al, wt.% | 6.5976 | 0.26121 | 6.0752 | 7.1201 | 5.814 | 7.3813 | 3.959% | 7.918% | 11.88% | 6.2677 | 6.9275
Ba, ppm | 4397.0592 | 464.1253 | 3468.8087 | 5325.3098 | 3004.6834 | 5789.4351 | 10.56% | 21.11% | 31.67% | 4177.2063 | 4616.9122
Be, ppm | 13.246 | 0.89319 | 11.4597 | 15.0324 | 10.5665 | 15.9256 | 6.743% | 13.49% | 20.23% | 12.5837 | 13.9083
Bi, ppm | 16.9218 | 1.0416 | 14.8385 | 19.005 | 13.7969 | 20.0466 | 6.155% | 12.31% | 18.47% | 16.0757 | 17.7679
Ca, wt.% | 0.89974 | 0.059978 | 0.77978 | 1.0197 | 0.7198 | 1.0797 | 6.666% | 13.33% | 20% | 0.85475 | 0.94472
Ce, wt.% | 3.9525 | 0.13505 | 3.6824 | 4.2226 | 3.5473 | 4.3577 | 3.417% | 6.834% | 10.25% | 3.7549 | 4.1501
Cr, ppm | 544.1822 | 41.9089 | 460.3644 | 628.0001 | 418.4554 | 669.909 | 7.701% | 15.4% | 23.1% | 516.9731 | 571.3913
Cs, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Dy, ppm | 216.7238 | 12.5779 | 191.5679 | 241.8797 | 178.99 | 254.4577 | 5.804% | 11.61% | 17.41% | 205.8876 | 227.56
Er, ppm | 50.459 | 3.1388 | 44.1814 | 56.7365 | 41.0426 | 59.8753 | 6.22% | 12.44% | 18.66% | 47.936 | 52.9819
Eu, ppm | 285.6091 | 11.4732 | 262.6626 | 308.5555 | 251.1893 | 320.0288 | 4.017% | 8.034% | 12.05% | 271.3286 | 299.8895
Fe, wt.% | 34.7122 | 0.63354 | 33.4451 | 35.9793 | 32.8116 | 36.6128 | 1.825% | 3.65% | 5.475% | 32.9766 | 36.4478
Ga, ppm | 213.7519 | 42.8101 | 128.1316 | 299.3721 | 85.3215 | 342.1822 | 20.03% | 40.06% | 60.08% | 203.0643 | 224.4395
Gd, ppm | 584.393 | 30.6922 | 523.0087 | 645.7774 | 492.3165 | 676.4696 | 5.252% | 10.5% | 15.76% | 555.1734 | 613.6127
Hf, ppm | 41.4062 | 7.2079 | 26.9904 | 55.8221 | 19.7825 | 63.03 | 17.41% | 34.82% | 52.22% | 39.3359 | 43.4766
Ho, ppm | 27.7013 | 2.1321 | 23.4371 | 31.9655 | 21.305 | 34.0977 | 7.697% | 15.39% | 23.09% | 26.3162 | 29.0864
In, ppm | 3.4687 | 0.22668 | 3.0153 | 3.9221 | 2.7887 | 4.1487 | 6.535% | 13.07% | 19.6% | 3.2953 | 3.6421
La, wt.% | 2.4128 | 0.082359 | 2.2481 | 2.5776 | 2.1658 | 2.6599 | 3.413% | 6.827% | 10.24% | 2.2922 | 2.5335
Lu, ppm | 2.1042 | 0.12164 | 1.8609 | 2.3474 | 1.7393 | 2.4691 | 5.781% | 11.56% | 17.34% | 1.999 | 2.2094
Mg, wt.% | 0.39174 | 0.020505 | 0.35073 | 0.43275 | 0.33023 | 0.45326 | 5.234% | 10.47% | 15.7% | 0.37216 | 0.41133
Mn, wt.% | 0.26252 | 0.022883 | 0.21676 | 0.30829 | 0.19387 | 0.33117 | 8.716% | 17.43% | 26.15% | 0.2494 | 0.27565
Mo, ppm | 113.945 | 5.6664 | 102.6123 | 125.2778 | 96.9459 | 130.9442 | 4.973% | 9.946% | 14.92% | 108.2478 | 119.6423
Nb, ppm | 4680.1729 | 301.2312 | 4077.7104 | 5282.6353 | 3776.4792 | 5583.8665 | 6.436% | 12.87% | 19.31% | 4446.1642 | 4914.1815
Nd, wt.% | 1.1764 | 0.050429 | 1.0755 | 1.2772 | 1.0251 | 1.3277 | 4.287% | 8.574% | 12.86% | 1.1176 | 1.2352
P, wt.% | 3.8122 | 0.12235 | 3.5675 | 4.0569 | 3.4451 | 4.1793 | 3.21% | 6.419% | 9.629% | 3.6216 | 4.0028
Pb, ppm | 506.1513 | 18.6321 | 468.8871 | 543.4156 | 450.2549 | 562.0477 | 3.681% | 7.362% | 11.04% | 480.8438 | 531.4589
Pr, ppm | 3771.8124 | 164.023 | 3443.7664 | 4099.8585 | 3279.7433 | 4263.8816 | 4.349% | 8.697% | 13.05% | 3583.2218 | 3960.4031
S, ppm | 1941 | 132.44 | 1676.1201 | 2205.8799 | 1543.6801 | 2338.3199 | 6.823% | 13.65% | 20.47% | 1843.95 | 2038.05
Si, wt.% | 1.5341 | 0.04334 | 1.4474 | 1.6207 | 1.404 | 1.6641 | 2.825% | 5.65% | 8.475% | 1.4574 | 1.6108
Sm, ppm | 1361.0146 | 36.1739 | 1288.6668 | 1433.3625 | 1252.4928 | 1469.5364 | 2.658% | 5.316% | 7.974% | 1292.9639 | 1429.0654
Sn, ppm | 135.8085 | 20.2242 | 95.3601 | 176.2569 | 75.1359 | 196.4811 | 14.89% | 29.78% | 44.68% | 129.0181 | 142.5989
Sr, ppm | 5203.8067 | 182.0822 | 4839.6423 | 5567.971 | 4657.5601 | 5750.0532 | 3.499% | 6.998% | 10.5% | 4943.6163 | 5463.997
Ta, ppm | 78.6692 | 5.0443 | 68.5806 | 88.7578 | 63.5363 | 93.8021 | 6.412% | 12.82% | 19.24% | 74.7357 | 82.6027
Tb, ppm | 56.5714 | 3.1201 | 50.3312 | 62.8116 | 47.2111 | 65.9317 | 5.515% | 11.03% | 16.55% | 53.7429 | 59.4
Th, ppm | 865.5289 | 42.8179 | 779.8931 | 951.1646 | 737.0753 | 993.9825 | 4.947% | 9.894% | 14.84% | 822.2524 | 908.8053
Ti, wt.% | 6.3017 | 0.18382 | 5.934 | 6.6693 | 5.7502 | 6.8531 | 2.917% | 5.834% | 8.751% | 5.9866 | 6.6168
Tm, ppm | 4.5221 | 0.26208 | 3.9979 | 5.0462 | 3.7358 | 5.3083 | 5.796% | 11.59% | 17.39% | 4.296 | 4.7482
U, ppm | 13.6006 | 0.40075 | 12.7991 | 14.4021 | 12.3983 | 14.8028 | 2.947% | 5.893% | 8.84% | 12.9205 | 14.2806
V, ppm | 534.1117 | 35.5843 | 462.9431 | 605.2804 | 427.3587 | 640.8647 | 6.662% | 13.32% | 19.99% | 507.4061 | 560.8173
W, ppm | 7.52 | 1.2382 | 5.0436 | 9.9964 | 3.8053 | 11.2347 | 16.47% | 32.93% | 49.4% | 7.144 | 7.896
Y, ppm | 523.8165 | 33.1896 | 457.4374 | 590.1957 | 424.2479 | 623.3852 | 6.336% | 12.67% | 19.01% | 497.6257 | 550.0074
Yb, ppm | 18.9818 | 0.72455 | 17.5327 | 20.4309 | 16.8081 | 21.1554 | 3.817% | 7.634% | 11.45% | 18.0327 | 19.9308
Zr, ppm | 1879.6006 | 203.1451 | 1473.3103 | 2285.8908 | 1270.1652 | 2489.036 | 10.81% | 21.62% | 32.42% | 1785.6205 | 1973.5806
4-Acid Digestion
Ag, ppm | 5.4832 | 0.35318 | 4.7768 | 6.1895 | 4.4236 | 6.5427 | 6.441% | 12.88% | 19.32% | 5.209 | 5.7573
Al, wt.% | 6.2092 | 0.39938 | 5.4104 | 7.0079 | 5.011 | 7.4073 | 6.432% | 12.86% | 19.3% | 5.8987 | 6.5196
Ba, ppm | 4359.1998 | 157.9633 | 4043.2733 | 4675.1263 | 3885.3101 | 4833.0896 | 3.624% | 7.247% | 10.87% | 4141.2398 | 4577.1598
Be, ppm | 11.6226 | 0.75763 | 10.1074 | 13.1379 | 9.3497 | 13.8955 | 6.519% | 13.04% | 19.56% | 11.0415 | 12.2038
Bi, ppm | 17.2678 | 1.0295 | 15.2089 | 19.3268 | 14.1794 | 20.3562 | 5.962% | 11.92% | 17.89% | 16.4044 | 18.1312
Ca, wt.% | 0.87181 | 0.061686 | 0.74844 | 0.99518 | 0.68675 | 1.0569 | 7.076% | 14.15% | 21.23% | 0.82822 | 0.9154
Cd, ppm | 1.203 | 0.14232 | 0.91842 | 1.4877 | 0.7761 | 1.63 | 11.83% | 23.66% | 35.49% | 1.1429 | 1.2632
Ce, wt.% | 3.9056 | 0.14877 | 3.6081 | 4.2032 | 3.4593 | 4.3519 | 3.809% | 7.618% | 11.43% | 3.7103 | 4.1009
Co, ppm | 18.7017 | 1.2612 | 16.1793 | 21.2241 | 14.9181 | 22.4853 | 6.744% | 13.49% | 20.23% | 17.7666 | 19.6368
Cs, ppm | < 0.1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Cu, ppm | 127.5667 | 5.3201 | 116.9265 | 138.2069 | 111.6063 | 143.527 | 4.17% | 8.341% | 12.51% | 121.1884 | 133.945
Dy, ppm | 215.1915 | 9.2587 | 196.6741 | 233.7088 | 187.4154 | 242.9675 | 4.303% | 8.605% | 12.91% | 204.4319 | 225.951
Er, ppm | 47.265 | 2.5083 | 42.2484 | 52.2816 | 39.7401 | 54.7899 | 5.307% | 10.61% | 15.92% | 44.9018 | 49.6283
Eu, ppm | 282.2001 | 11.8075 | 258.5852 | 305.8151 | 246.7778 | 317.6225 | 4.184% | 8.368% | 12.55% | 268.0901 | 296.3101
Fe, wt.% | 29.5527 | 2.0248 | 25.5031 | 33.6022 | 23.4784 | 35.627 | 6.851% | 13.7% | 20.55% | 28.075 | 31.0303
Ga, ppm | 188.3139 | 30.7797 | 126.7545 | 249.8734 | 95.9748 | 280.6531 | 16.34% | 32.69% | 49.03% | 178.8982 | 197.7296
Gd, ppm | 581.0732 | 35.8098 | 509.4535 | 652.6928 | 473.6437 | 688.5027 | 6.163% | 12.33% | 18.49% | 552.0195 | 610.1268
Hf, ppm | 14.4049 | 1.9332 | 10.5385 | 18.2713 | 8.6053 | 20.2045 | 13.42% | 26.84% | 40.26% | 13.6846 | 15.1251
Ho, ppm | 26.7772 | 0.93535 | 24.9065 | 28.6479 | 23.9712 | 29.5832 | 3.493% | 6.986% | 10.48% | 25.4383 | 28.116
In, ppm | 3.1831 | 0.20817 | 2.7668 | 3.5995 | 2.5586 | 3.8076 | 6.54% | 13.08% | 19.62% | 3.024 | 3.3423
La, wt.% | 2.2739 | 0.092966 | 2.088 | 2.4598 | 1.995 | 2.5528 | 4.088% | 8.177% | 12.27% | 2.1602 | 2.3876
Li, ppm | 3.0419 | 0.332 | 2.3779 | 3.7059 | 2.0459 | 4.0379 | 10.91% | 21.83% | 32.74% | 2.8898 | 3.1939
Lu, ppm | 1.7206 | 0.28169 | 1.1572 | 2.284 | 0.87556 | 2.5657 | 16.37% | 32.74% | 49.11% | 1.6346 | 1.8067
Mg, wt.% | 0.37383 | 0.0261 | 0.32163 | 0.42603 | 0.29553 | 0.45213 | 6.982% | 13.96% | 20.95% | 0.35514 | 0.39252
Mn, wt.% | 0.19771 | 0.022452 | 0.15281 | 0.24262 | 0.13036 | 0.26507 | 11.36% | 22.71% | 34.07% | 0.18783 | 0.2076
Mo, ppm | 98.0101 | 10.1481 | 77.714 | 118.3062 | 67.5659 | 128.4543 | 10.35% | 20.71% | 31.06% | 93.1096 | 102.9106
Na, wt.% | < 0.2 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Nd, wt.% | 1.0951 | 0.043014 | 1.0091 | 1.1812 | 0.96611 | 1.2242 | 3.928% | 7.855% | 11.78% | 1.0404 | 1.1499
Ni, ppm | 105.8075 | 13.5473 | 78.7129 | 132.9021 | 65.1656 | 146.4494 | 12.8% | 25.61% | 38.41% | 100.5171 | 111.0978
P, wt.% | 3.1508 | 0.37551 | 2.3998 | 3.9018 | 2.0243 | 4.2773 | 11.92% | 23.84% | 35.75% | 2.9933 | 3.3084
Pb, ppm | 572.5402 | 16.0717 | 540.3969 | 604.6835 | 524.3253 | 620.7552 | 2.807% | 5.614% | 8.421% | 543.9132 | 601.1672
Pr, ppm | 3669.9075 | 192.7586 | 3284.3903 | 4055.4246 | 3091.6318 | 4248.1832 | 5.252% | 10.5% | 15.76% | 3486.4121 | 3853.4029
Rb, ppm | 0.42577 | 0.054292 | 0.31718 | 0.53435 | 0.26289 | 0.58864 | 12.75% | 25.5% | 38.25% | 0.40448 | 0.44706
Re, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Sc, ppm | 148.5346 | 9.0796 | 130.3755 | 166.6938 | 121.296 | 175.7733 | 6.113% | 12.23% | 18.34% | 141.1079 | 155.9614
Sm, ppm | 1306.679 | 80.8176 | 1145.0439 | 1468.3142 | 1064.2263 | 1549.1318 | 6.185% | 12.37% | 18.55% | 1241.3451 | 1372.013
Sr, wt.% | 0.50533 | 0.02625 | 0.45283 | 0.55783 | 0.42658 | 0.58408 | 5.195% | 10.39% | 15.58% | 0.48006 | 0.53059
Tb, ppm | 57.0775 | 3.3289 | 50.4197 | 63.7352 | 47.0908 | 67.0641 | 5.832% | 11.66% | 17.5% | 54.2236 | 59.9313
Te, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Th, ppm | 804.6092 | 149.6429 | 505.3233 | 1103.8951 | 355.6804 | 1253.538 | 18.6% | 37.2% | 55.79% | 764.3787 | 844.8397
Tl, ppm | 0.08732 | 0.0089241 | 0.069472 | 0.10517 | 0.060548 | 0.11409 | 10.22% | 20.44% | 30.66% | 0.082954 | 0.091686
Tm, ppm | 3.8238 | 0.2742 | 3.2754 | 4.3722 | 3.0012 | 4.6464 | 7.171% | 14.34% | 21.51% | 3.6326 | 4.015
U, ppm | 12.6051 | 0.67854 | 11.248 | 13.9622 | 10.5695 | 14.6408 | 5.383% | 10.77% | 16.15% | 11.9749 | 13.2354
V, ppm | 426.604 | 68.9831 | 288.6379 | 564.5702 | 219.6549 | 633.5532 | 16.17% | 32.34% | 48.51% | 405.2738 | 447.9342
Y, ppm | 477.799 | 32.3574 | 413.0842 | 542.5138 | 380.7268 | 574.8713 | 6.772% | 13.54% | 20.32% | 453.9091 | 501.689
Yb, ppm | 14.9004 | 0.95238 | 12.9956 | 16.8052 | 12.0433 | 17.7576 | 6.392% | 12.78% | 19.17% | 14.1554 | 15.6454
Zn, ppm | 921.4187 | 133.0092 | 655.4003 | 1187.4371 | 522.3912 | 1320.4463 | 14.44% | 28.87% | 43.31% | 875.3478 | 967.4896

## Fusion XRF
| Analytical results for Al in OREAS 465 (Indicative Value 6.68 wt.%) | [8 empty] | Indicative
Al | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 6.653 | 6.563 | 6.721 | 6.774 |  |  |  | 1
| 1 | 2 | 6.668 | 6.616 | 6.721 | 6.721 |  |  |  | 1
| 1 | 3 | 6.679 | 6.51 | 6.668 | 6.774 |  |  |  | 16
| 1 | 4 | 6.705 | 6.616 | 6.668 | 6.774 |  |  |  | 6.6786
| 1 | 5 | 6.705 | 6.563 | 6.721 | 6.721 |  |  |  | 7
| 1 | 6 | 6.684 | 6.616 | 6.668 | 6.774
| Mean |  | 6.6823 | 6.5807 | 6.6945 | 6.7563
| Median |  | 6.6815 | 6.5895 | 6.6945 | 6.774
| Std Dev. |  | 0.02053 | 0.043274 | 0.029029 | 0.027369
| Rel.Std.Dev. |  | 0.3072% | 0.6576% | 0.4336% | 0.4051%
| PDM3 |  | 0.05628% | -1.466% | 0.2385% | 1.164%
| Z-Score (Absolute) |  | 0.11 | 1.97 | 0.11 | 1.24
| Analytical results for Ba in OREAS 465 (Indicative Value 4784 ppm) | [8 empty] | Indicative
Ba | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 5020 | 4750 | [5 empty] | 1
| 1 | 2 | 4840 | 4660 | [5 empty] | 2
| 1 | 3 | 4840 | 4750 | [5 empty] | 16
| 1 | 4 | 4840 | 4570 | [5 empty] | 4784.2962
| 1 | 5 | 4840 | 4840 | [5 empty] | 8
| 1 | 6 | 4840 | 4660
| Mean |  | 4870 | 4705
| Median |  | 4840 | 4705
| Std Dev. |  | 73.4847 | 94.3928
| Rel.Std.Dev. |  | 1.509% | 2.006%
| PDM3 |  | 1.791% | -1.657%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ca in OREAS 465 (Indicative Value 0.938 wt.%) | [8 empty] | Indicative
Ca | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.922 | 0.929 | 0.929 | 0.958 |  |  |  | 1
| 1 | 2 | 0.915 | 0.958 | 0.943 | 0.965 |  |  |  | 3
| 1 | 3 | 0.922 | 0.936 | 0.915 | 0.965 |  |  |  | 16
| 1 | 4 | 0.915 | 0.943 | 0.901 | 0.951 |  |  |  | 0.93804
| 1 | 5 | 0.922 | 0.943 | 0.943 | 0.965 |  |  |  | 9
| 1 | 6 | 0.922 | 0.951 | 0.929 | 0.972
| Mean |  | 0.91967 | 0.94333 | 0.92667 | 0.96267
| Median |  | 0.922 | 0.943 | 0.929 | 0.965
| Std Dev. |  | 0.0036148 | 0.010328 | 0.016367 | 0.0072296
| Rel.Std.Dev. |  | 0.3931% | 1.095% | 1.766% | 0.751%
| PDM3 |  | -1.958% | 0.5647% | -1.212% | 2.626%
| Z-Score (Absolute) |  | 0.88 | 0.47 | 0.47 | 1.55
| Analytical results for Ce in OREAS 465 (Certified Value 3.97 wt.%) | [8 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 4.03 | 3.973 | 3.791 | 3.997 | 4.098 | 3.97 | 3.932 | 1
| 1 | 2 | 4.04 | 3.956 | 4.004 | 3.965 | 4.115 | 3.99 | 3.94 | 0
| 1 | 3 | 4 | 3.948 | 3.876 | 3.973 | 4.115 | 3.99 | 3.924 | 16
| 1 | 4 | 4.04 | 3.973 | 3.987 | 3.924 | 4.098 | 3.98 | 3.932 | 3.9692
| 1 | 5 | 4 | 3.973 | 3.91 | 3.989 | 4.098 | 3.99 | 3.924 | 7
| 1 | 6 | 4.01 | 3.973 | 3.97 | 3.948 | 4.115 | 3.98 | 3.932
| Mean |  | 4.02 | 3.966 | 3.923 | 3.966 | 4.1065 | 3.9833 | 3.9307
| Median |  | 4.02 | 3.973 | 3.94 | 3.969 | 4.1065 | 3.985 | 3.932
| Std Dev. |  | 0.018974 | 0.011136 | 0.080831 | 0.026937 | 0.0093113 | 0.008165 | 0.0060222
| Rel.Std.Dev. |  | 0.472% | 0.2808% | 2.06% | 0.6792% | 0.2267% | 0.205% | 0.1532%
| PDM3 |  | 1.28% | -0.08022% | -1.164% | -0.08022% | 3.46% | 0.3565% | -0.9704%
| Z-Score (Absolute) |  | 2.09 | 0 | 0.63 | 0 | 5.43 | 0.67 | 1.36
| Analytical results for Cr in OREAS 465 (Indicative Value 503 ppm) | [8 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 680 | 376 | [5 empty] | 1
| 1 | 2 | 620 | 356 | [5 empty] | 4
| 1 | 3 | 680 | 383 | [5 empty] | 16
| 1 | 4 | 620 | 308 | [5 empty] | 502.908
| 1 | 5 | 680 | 342 | [5 empty] | 10
| 1 | 6 | 620 | 369
| Mean |  | 650 | 355.6667
| Median |  | 650 | 362.5
| Std Dev. |  | 32.8634 | 27.5584
| Rel.Std.Dev. |  | 5.056% | 7.748%
| PDM3 |  | 29.25% | -29.28%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Dy in OREAS 465 (Indicative Value 287 ppm) | [8 empty] | Indicative
Dy | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 260 | 350 | 253 |  |  |  |  | 1
| 1 | 2 | 260 | 350 | 253 |  |  |  |  | 5
| 1 | 3 | 260 | 350 | 253 |  |  |  |  | 16
| 1 | 4 | 260 | 350 | 253 |  |  |  |  | 287.0675
| 1 | 5 | 260 | 350 | 253 |  |  |  |  | 11
| 1 | 6 | 260 | 350 | 253
| Mean |  | 260 | 350 | 253
| Median |  | 260 | 350 | 253
| Std Dev. |  | 0 | 0 | 0
| Rel.Std.Dev. |  | 0% | 0% | 0%
| PDM3 |  | -9.429% | 21.92% | -11.87%
| Z-Score (Absolute) |  | 0 | 8.15 | 0.67
| Analytical results for Er in OREAS 465 (Indicative Value 54 ppm) | [8 empty] | Indicative
Er | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | < 90 | 90 | 35 |  |  |  |  | 1
| 1 | 2 | < 90 | 90 | 35 |  |  |  |  | 6
| 1 | 3 | < 90 | < 90 | 35 |  |  |  |  | 16
| 1 | 4 | < 90 | 90 | 35 |  |  |  |  | 53.928
| 1 | 5 | < 90 | 90 | 35 |  |  |  |  | 12
| 1 | 6 | < 90 | < 90 | 35
| Mean |  |  | 90 | 35
| Median |  |  | 90 | 35
| Std Dev. |  |  | 0 | 0
| Rel.Std.Dev. |  |  | 0% | 0%
| PDM3 |  |  | 66.89% | -35.1%
| Z-Score (Absolute) |  | NA | 0.67 | 0.67
| Analytical results for Eu in OREAS 465 (Indicative Value 302 ppm) | [8 empty] | Indicative
Eu | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 09 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 430 | 260 | 260 |  |  |  |  | 1
| 1 | 2 | 350 | 260 | 260 |  |  |  |  | 7
| 1 | 3 | 430 | 260 | 260 |  |  |  |  | 16
| 1 | 4 | 430 | 260 | 260 |  |  |  |  | 302.2714
| 1 | 5 | 350 | 260 | 260 |  |  |  |  | 13
| 1 | 6 | 350 | 260 | 260
| Mean |  | 390 | 260 | 260
| Median |  | 390 | 260 | 260
| Std Dev. |  | 43.8178 | 0 | 0
| Rel.Std.Dev. |  | 11.24% | 0% | 0%
| PDM3 |  | 29.02% | -13.98% | -13.98%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Fe in OREAS 465 (Certified Value 34.94 wt.%) | [8 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 09 | 13 | 14 | 15 | 22 |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  | 2
3
| 1 | 1 | 35 | 34.972 | 34.55 | 33.573 | 36.09 |  |  | 1
| 1 | 2 | 34.937 | 36.231 | 34.41 | 33.713 | 36.09 |  |  | 0
| 1 | 3 | 34.909 | 35.672 | 34.27 | 33.643 | 35.88 |  |  | 16
| 1 | 4 | 35.063 | 35.952 | 34.13 | 33.643 | 35.88 |  |  | 34.9442
| 1 | 5 | 34.979 | 35.462 | 34.48 | 33.573 | 36.09 |  |  | 8
| 1 | 6 | 34.937 | 36.091 | 34.34 | 33.643 | 36.09
| Mean |  | 34.9708 | 35.73 | 34.3633 | 33.6313 | 36.02
| Median |  | 34.958 | 35.812 | 34.375 | 33.643 | 36.09
| Std Dev. |  | 0.055722 | 0.46484 | 0.15122 | 0.052694 | 0.10844
| Rel.Std.Dev. |  | 0.1593% | 1.301% | 0.4401% | 0.1567% | 0.3011%
| PDM3 |  | 0.07633% | 2.249% | -1.662% | -3.757% | 3.079%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.54 | 1.19 | 0.93
| Analytical results for Gd in OREAS 465 (Indicative Value 755 ppm) | [8 empty] | Indicative
Gd | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 09 | 14 | 22 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | 780 | 870 | 690 | 690 |  |  |  | 1
| 1 | 2 | 780 | 870 | 690 | 690 |  |  |  | 8
| 1 | 3 | 780 | 870 | 690 | 690 |  |  |  | 16
| 1 | 4 | 690 | 870 | 780 | 690 |  |  |  | 755.0574
| 1 | 5 | 690 | 870 | 690 | 690 |  |  |  | 14
| 1 | 6 | 780 | 870 | 690 | 690
| Mean |  | 750 | 870 | 705 | 690
| Median |  | 780 | 870 | 690 | 690
| Std Dev. |  | 46.4758 | 0 | 36.7423 | 0
| Rel.Std.Dev. |  | 6.197% | 0% | 5.212% | 0%
| PDM3 |  | -0.6698% | 15.22% | -6.63% | -8.616%
| Z-Score (Absolute) |  | 0.5 | 3.34 | 0.5 | 0.85
| Analytical results for Hf in OREAS 465 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Hf | Round | Replicate | Lab | [6 empty] | 1
| No | No | 13 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 9
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 15
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 465 (Indicative Value 26.2 ppm) | [8 empty] | Indicative
Ho | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | 26 | [5 empty] | 1
| 1 | 2 | < 90 | 26 | [5 empty] | 10
| 1 | 3 | < 90 | 26 | [5 empty] | 16
| 1 | 4 | < 90 | 26 | [5 empty] | 26.1894
| 1 | 5 | < 90 | 26 | [5 empty] | 16
| 1 | 6 | < 90 | 26
| Mean |  |  | 26
| Median |  |  | 26
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | -0.7233%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for K in OREAS 465 (Indicative Value 0.012 wt.%) | [8 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.017 | < 0.008 | 0.008 | 0.017 |  |  |  | 1
| 1 | 2 | 0.017 | < 0.008 | 0.008 | 0.017 |  |  |  | 11
| 1 | 3 | 0.017 | < 0.008 | < 0.008 | 0.017 |  |  |  | 16
| 1 | 4 | 0.017 | 0.008 | 0.025 | 0.017 |  |  |  | 0.012279
| 1 | 5 | 0.017 | < 0.008 | 0.017 | 0.017 |  |  |  | 17
| 1 | 6 | 0.017 | < 0.008 | 0.008 | 0.017
| Mean |  | 0.017 | 0.008 | 0.0132 | 0.017
| Median |  | 0.017 | 0.008 | 0.008 | 0.017
| Std Dev. |  | 0 |  | 0.0076616 | 0
| Rel.Std.Dev. |  | 0% |  | 58.04% | 0%
| PDM3 |  | 38.44% | -34.85% | 7.498% | 38.44%
| Z-Score (Absolute) |  | 0.67 | 2.79 | 0.67 | 0.67
| Analytical results for La in OREAS 465 (Certified Value 2.42 wt.%) | [8 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 2.413 | 2.387 | 2.353 | 2.405 | 2.439 | 2.47 | 2.405 | 1
| 1 | 2 | 2.422 | 2.396 | 2.464 | 2.422 | 2.447 | 2.47 | 2.413 | 0
| 1 | 3 | 2.405 | 2.387 | 2.422 | 2.405 | 2.439 | 2.47 | 2.396 | 16
| 1 | 4 | 2.405 | 2.396 | 2.413 | 2.439 | 2.439 | 2.47 | 2.405 | 2.4205
| 1 | 5 | 2.405 | 2.396 | 2.405 | 2.405 | 2.43 | 2.46 | 2.43 | 9
| 1 | 6 | 2.396 | 2.387 | 2.43 | 2.405 | 2.447 | 2.47 | 2.405
| Mean |  | 2.4077 | 2.3915 | 2.4145 | 2.4135 | 2.4402 | 2.4683 | 2.409
| Median |  | 2.405 | 2.3915 | 2.4175 | 2.405 | 2.439 | 2.47 | 2.405
| Std Dev. |  | 0.0088468 | 0.0049295 | 0.036391 | 0.014223 | 0.0063377 | 0.0040825 | 0.01161
| Rel.Std.Dev. |  | 0.3674% | 0.2061% | 1.507% | 0.5893% | 0.2597% | 0.1654% | 0.482%
| PDM3 |  | -0.5312% | -1.199% | -0.2488% | -0.2902% | 0.8115% | 1.975% | -0.4761%
| Z-Score (Absolute) |  | 0.67 | 2.53 | 0.17 | 0 | 3.2 | 6.56 | 0.51
| Analytical results for Lu in OREAS 465 (Indicative Value < 20 ppm) | [8 empty] | Indicative
Lu | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 18 | [5 empty] | 1
| 1 | 2 | < 90 | < 18 | [5 empty] | 1
| 1 | 3 | < 90 | < 18 | [5 empty] | 16
| 1 | 4 | < 90 | < 18 | [5 empty] | < 20
| 1 | 5 | < 90 | < 18 | [5 empty] | 7
| 1 | 6 | < 90 | < 18
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Mg in OREAS 465 (Indicative Value 0.4 wt.%) | [8 empty] | Indicative
Mg | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.416 | 0.271 | 0.404 | 0.398 |  |  |  | 1
| 1 | 2 | 0.416 | 0.41 | 0.386 | 0.38 |  |  |  | 2
| 1 | 3 | 0.416 | 0.404 | 0.392 | 0.392 |  |  |  | 16
| 1 | 4 | 0.416 | 0.398 | 0.386 | 0.398 |  |  |  | 0.40018
| 1 | 5 | 0.41 | 0.398 | 0.374 | 0.398 |  |  |  | 8
| 1 | 6 | 0.41 | 0.398 | 0.398 | 0.404
| Mean |  | 0.414 | 0.37983 | 0.39 | 0.395
| Median |  | 0.416 | 0.398 | 0.389 | 0.398
| Std Dev. |  | 0.0030984 | 0.053533 | 0.010507 | 0.0082704
| Rel.Std.Dev. |  | 0.7484% | 14.09% | 2.694% | 2.094%
| PDM3 |  | 3.453% | -5.085% | -2.545% | -1.295%
| Z-Score (Absolute) |  | 1.83 | 0.38 | 0.96 | 0.38
| Analytical results for Mn in OREAS 465 (Indicative Value 0.286 wt.%) | [8 empty] | Indicative
Mn | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.279 | 0.294 | 0.287 | 0.271 |  |  |  | 1
| 1 | 2 | 0.279 | 0.31 | 0.287 | 0.279 |  |  |  | 3
| 1 | 3 | 0.287 | 0.302 | 0.279 | 0.279 |  |  |  | 16
| 1 | 4 | 0.279 | 0.31 | 0.279 | 0.279 |  |  |  | 0.28559
| 1 | 5 | 0.279 | 0.294 | 0.287 | 0.279 |  |  |  | 9
| 1 | 6 | 0.279 | 0.294 | 0.287 | 0.279
| Mean |  | 0.28033 | 0.30067 | 0.28433 | 0.27767
| Median |  | 0.279 | 0.298 | 0.287 | 0.279
| Std Dev. |  | 0.003266 | 0.0078655 | 0.0041312 | 0.003266
| Rel.Std.Dev. |  | 1.165% | 2.616% | 1.453% | 1.176%
| PDM3 |  | -1.84% | 5.28% | -0.439% | -2.773%
| Z-Score (Absolute) |  | 0.4 | 3.91 | 0.4 | 0.94
| Analytical results for Na in OREAS 465 (Indicative Value 0.128 wt.%) | [8 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 3
3
| 1 | 1 | 0.126 | 0.415 | < 0.07 | 0.104 |  |  |  | 1
| 1 | 2 | 0.119 | 0.171 | < 0.07 | 0.134 |  |  |  | 4
| 1 | 3 | 0.111 | 0.156 | < 0.07 | 0.111 |  |  |  | 16
| 1 | 4 | 0.104 | 0.126 | < 0.07 | 0.119 |  |  |  | 0.12751
| 1 | 5 | 0.111 | 0.148 | < 0.07 | 0.119 |  |  |  | 10
| 1 | 6 | 0.104 | 0.156 | < 0.07 | 0.126
| Mean |  | 0.1125 | 0.19533 |  | 0.11883
| Median |  | 0.111 | 0.156 |  | 0.119
| Std Dev. |  | 0.0086429 | 0.10862 |  | 0.01061
| Rel.Std.Dev. |  | 7.683% | 55.61% |  | 8.928%
| PDM3 |  | -11.77% | 53.19% |  | -6.808%
| Z-Score (Absolute) |  | 0.67 | 3.56 | NA | 0
| Analytical results for Nb in OREAS 465 (Indicative Value 4649 ppm) | [8 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 14 | 15 | 16 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 5100 | 4680 | 4700 |  |  |  |  | 1
| 1 | 2 | 5100 | 4680 | 4600 |  |  |  |  | 27
| 1 | 3 | 5030 | 4680 | 4700 |  |  |  |  | 16
| 1 | 4 | 5030 | 4540 | 4600 |  |  |  |  | 4649.3621
| 1 | 5 | 5100 | 4680 | 4700 |  |  |  |  | 11
| 1 | 6 | 5100 | 4610 | 4600
| Mean |  | 5076.6667 | 4645 | 4650
| Median |  | 5100 | 4680 | 4650
| Std Dev. |  | 36.1478 | 58.5662 | 54.7723
| Rel.Std.Dev. |  | 0.712% | 1.261% | 1.178%
| PDM3 |  | 9.191% | -0.09382% | 0.01372%
| Z-Score (Absolute) |  | 227.17 | 0.67 | 0
| Analytical results for Nd in OREAS 465 (Certified Value 1.18 wt.%) | [8 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 2
3
| 1 | 1 | 1.183 | 1.149 | 0.995 | 1.295 | 1.269 | 1.17 | 1.1403 | 1
| 1 | 2 | 1.2 | 1.149 | 1.046 | 1.303 | 1.286 | 1.17 | 1.1403 | 0
| 1 | 3 | 1.183 | 1.149 | 1.029 | 1.32 | 1.295 | 1.16 | 1.1403 | 16
| 1 | 4 | 1.2 | 1.149 | 1.037 | 1.295 | 1.286 | 1.16 | 1.1231 | 1.1807
| 1 | 5 | 1.183 | 1.149 | 1.02 | 1.303 | 1.295 | 1.17 | 1.1403 | 10
| 1 | 6 | 1.183 | 1.149 | 1.072 | 1.295 | 1.312 | 1.16 | 1.1403
| Mean |  | 1.1887 | 1.149 | 1.0332 | 1.3018 | 1.2905 | 1.165 | 1.1374
| Median |  | 1.183 | 1.149 | 1.033 | 1.299 | 1.2905 | 1.165 | 1.1403
| Std Dev. |  | 0.0087788 | 0 | 0.025826 | 0.0097245 | 0.014181 | 0.0054772 | 0.0070219
| Rel.Std.Dev. |  | 0.7385% | 0% | 2.5% | 0.747% | 1.099% | 0.4701% | 0.6173%
| PDM3 |  | 0.6712% | -2.688% | -12.5% | 10.26% | 9.296% | -1.333% | -3.668%
| Z-Score (Absolute) |  | 0.58 | 0.39 | 3.22 | 3.34 | 3.06 | 0 | 0.67
| Analytical results for Ni in OREAS 465 (Indicative Value 157 ppm) | [8 empty] | Indicative
Ni | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 160 | [6 empty] | 1
| 1 | 2 | 160 | [6 empty] | 6
| 1 | 3 | 160 | [6 empty] | 16
| 1 | 4 | 160 | [6 empty] | 157.1709
| 1 | 5 | 160 | [6 empty] | 12
| 1 | 6 | 160
| Mean |  | 160
| Median |  | 160
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 1.8%
| Z-Score (Absolute) |  | NA
| Analytical results for P in OREAS 465 (Indicative Value 3.83 wt.%) | [8 empty] | Indicative
P | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 09 | 14 | 15 |  |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 2
2
| 1 | 1 | 3.7899 | 3.805 | 3.888 |  |  |  |  | 1
| 1 | 2 | 3.7895 | 3.81 | 3.923 |  |  |  |  | 7
| 1 | 3 | 3.772 | 3.77 | 3.901 |  |  |  |  | 16
| 1 | 4 | 3.7917 | 3.788 | 3.897 |  |  |  |  | 3.8275
| 1 | 5 | 3.7812 | 3.81 | 3.884 |  |  |  |  | 13
| 1 | 6 | 3.7847 | 3.783 | 3.927
| Mean |  | 3.7848 | 3.7943 | 3.9033
| Median |  | 3.7871 | 3.7965 | 3.899
| Std Dev. |  | 0.0073826 | 0.016525 | 0.017896
| Rel.Std.Dev. |  | 0.1951% | 0.4355% | 0.4585%
| PDM3 |  | -1.115% | -0.8667% | 1.981%
| Z-Score (Absolute) |  | 0.67 | 0 | 7.78
| Analytical results for Pb in OREAS 465 (Indicative Value 603 ppm) | [8 empty] | Indicative
Pb | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 650 | [6 empty] | 1
| 1 | 2 | 650 | [6 empty] | 8
| 1 | 3 | 560 | [6 empty] | 16
| 1 | 4 | 560 | [6 empty] | 603.4163
| 1 | 5 | 560 | [6 empty] | 14
| 1 | 6 | 650
| Mean |  | 605
| Median |  | 605
| Std Dev. |  | 49.295
| Rel.Std.Dev. |  | 8.148%
| PDM3 |  | 0.2625%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 465 (Certified Value 3752 ppm) | [8 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 3310 | 3810 | 4020 | 3890 | 4020 | 3700 | 3640 | 1
| 1 | 2 | 3390 | 3720 | 4530 | 3890 | 3930 | 3700 | 3560 | 0
| 1 | 3 | 3310 | 3720 | 4530 | 3720 | 3760 | 3700 | 3560 | 16
| 1 | 4 | 3390 | 3810 | 3930 | 3810 | 3850 | 3600 | 3560 | 3752.3079
| 1 | 5 | 3390 | 3810 | 4100 | 3720 | 3930 | 3700 | 3560 | 11
| 1 | 6 | 3390 | 3810 | 4190 | 3810 | 3850 | 3700 | 3640
| Mean |  | 3363.3333 | 3780 | 4216.6667 | 3806.6667 | 3890 | 3683.3333 | 3586.6667
| Median |  | 3390 | 3810 | 4145 | 3810 | 3890 | 3700 | 3560
| Std Dev. |  | 41.3118 | 46.4758 | 257.5008 | 76.0701 | 89.666 | 40.8248 | 41.3118
| Rel.Std.Dev. |  | 1.228% | 1.23% | 6.107% | 1.998% | 2.305% | 1.108% | 1.152%
| PDM3 |  | -10.37% | 0.738% | 12.38% | 1.449% | 3.67% | -1.838% | -4.414%
| Z-Score (Absolute) |  | 2.58 | 0 | 2.72 | 0.17 | 0.67 | 0.5 | 1.2
| Analytical results for Si in OREAS 465 (Indicative Value 1.5 wt.%) | [8 empty] | Indicative
Si | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 1.444 | 1.342 | 1.529 | 1.524 |  |  |  | 1
| 1 | 2 | 1.44 | 1.529 | 1.659 | 1.51 |  |  |  | 9
| 1 | 3 | 1.458 | 1.482 | 1.491 | 1.524 |  |  |  | 16
| 1 | 4 | 1.458 | 1.51 | 1.533 | 1.529 |  |  |  | 1.4973
| 1 | 5 | 1.458 | 1.463 | 1.547 | 1.515 |  |  |  | 15
| 1 | 6 | 1.435 | 1.505 | 1.519 | 1.51
| Mean |  | 1.4488 | 1.4718 | 1.5463 | 1.5187
| Median |  | 1.451 | 1.4935 | 1.531 | 1.5195
| Std Dev. |  | 0.010439 | 0.067621 | 0.058277 | 0.0080911
| Rel.Std.Dev. |  | 0.7205% | 4.594% | 3.769% | 0.5328%
| PDM3 |  | -3.239% | -1.703% | 3.273% | 1.425%
| Z-Score (Absolute) |  | 3.04 | 0.53 | 0.81 | 0.53
| Analytical results for Sm in OREAS 465 (Certified Value 1414 ppm) | [8 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 1380 | 1380 | 4140 | 1120 | 1380 | 1600 | 1470 | 1
| 1 | 2 | 1380 | 1380 | 4310 | 1210 | 1290 | 1600 | 1470 | 0
| 1 | 3 | 1380 | 1470 | - | 1290 | 1290 | 1600 | 1470 | 16
| 1 | 4 | 1380 | 1470 | > 4300 | 1290 | 1380 | 1600 | 1470 | 1414.1286
| 1 | 5 | 1380 | 1380 | 4310 | 1290 | 1550 | 1600 | 1470 | 12
| 1 | 6 | 1380 | 1380 | 4230 | 1290 | 1380 | 1600 | 1470
| Mean |  | 1380 | 1410 | 4247.5 | 1248.3333 | 1378.3333 | 1600 | 1470
| Median |  | 1380 | 1380 | 4270 | 1290 | 1380 | 1600 | 1470
| Std Dev. |  | 0 | 46.4758 | 80.9835 | 70.5455 | 94.9561 | 2.4908e-13 | 0
| Rel.Std.Dev. |  | 0% | 3.296% | 1.907% | 5.651% | 6.889% | 1.557e-14% | 0%
| PDM3 |  | -2.413% | -0.292% | 200.4% | -11.72% | -2.531% | 13.14% | 3.951%
| Z-Score (Absolute) |  | 0.33 | 0 | 33.29 | 1.85 | 0.34 | 2.25 | 0.67
| Analytical results for Sn in OREAS 465 (Indicative Value 171 ppm) | [8 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 240 | 160 | [5 empty] | 1
| 1 | 2 | 160 | 80 | [5 empty] | 10
| 1 | 3 | 240 | 160 | [5 empty] | 16
| 1 | 4 | 240 | 160 | [5 empty] | 170.6574
| 1 | 5 | 240 | 80 | [5 empty] | 16
| 1 | 6 | 240 | 80
| Mean |  | 226.6667 | 120
| Median |  | 240 | 120
| Std Dev. |  | 32.6599 | 43.8178
| Rel.Std.Dev. |  | 14.41% | 36.51%
| PDM3 |  | 32.82% | -29.68%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for S in OREAS 465 (Indicative Value 0.194 wt.%) | [8 empty] | Indicative
S | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | BF*XRF | [6 empty] | 3
3
| 1 | 1 | 0.2018 | [6 empty] | 1
| 1 | 2 | 0.2014 | [6 empty] | 11
| 1 | 3 | 0.2002 | [6 empty] | 16
| 1 | 4 | 0.1862 | [6 empty] | 0.19422
| 1 | 5 | 0.1906 | [6 empty] | 17
| 1 | 6 | 0.185
| Mean |  | 0.1942
| Median |  | 0.1954
| Std Dev. |  | 0.0078384
| Rel.Std.Dev. |  | 4.036%
| PDM3 |  | -0.009023%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 465 (Indicative Value 0.51 wt.%) | [8 empty] | Indicative
Sr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | wt.%
|  |  | BF*XRF | BF*XRF | [5 empty] | 3
3
| 1 | 1 | 0.541 | 0.465 | [5 empty] | 1
| 1 | 2 | 0.541 | 0.482 | [5 empty] | 1
| 1 | 3 | 0.541 | 0.482 | [5 empty] | 16
| 1 | 4 | 0.541 | 0.474 | [5 empty] | 0.51018
| 1 | 5 | 0.55 | 0.474 | [5 empty] | 7
| 1 | 6 | 0.541 | 0.49
| Mean |  | 0.5425 | 0.47783
| Median |  | 0.541 | 0.478
| Std Dev. |  | 0.0036742 | 0.0086814
| Rel.Std.Dev. |  | 0.6773% | 1.817%
| PDM3 |  | 6.336% | -6.339%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Ta in OREAS 465 (Indicative Value < 100 ppm) | [8 empty] | Indicative
Ta | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 27
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 8
| 1 | 6 | < 80
| Mean |  | 80
| Median |  | 80
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 465 (Indicative Value 57 ppm) | [8 empty] | Indicative
Tb | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | < 90 | 51 | [5 empty] | 1
| 1 | 2 | < 90 | 68 | [5 empty] | 3
| 1 | 3 | < 90 | 51 | [5 empty] | 16
| 1 | 4 | < 90 | 51 | [5 empty] | 56.6797
| 1 | 5 | < 90 | 51 | [5 empty] | 9
| 1 | 6 | < 90 | 68
| Mean |  |  | 56.6667
| Median |  |  | 51
| Std Dev. |  |  | 8.7788
| Rel.Std.Dev. |  |  | 15.49%
| PDM3 |  |  | -0.023%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 465 (Certified Value 792 ppm) | [8 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  | 0
0
| 1 | 1 | 790 | 880 | 620 | 700 | 967 | 820 |  | 1
| 1 | 2 | 790 | 880 | 620 | 700 | 967 | 820 |  | 0
| 1 | 3 | 780 | 880 | 620 | 700 | 967 | 810 |  | 16
| 1 | 4 | 780 | 880 | 620 | 700 | 879 | 810 |  | 792.0682
| 1 | 5 | 800 | 880 | 620 | 700 | 967 | 800 |  | 13
| 1 | 6 | 790 | 880 | 620 | 700 | 967 | 830
| Mean |  | 788.3333 | 880 | 620 | 700 | 952.3333 | 815
| Median |  | 790 | 880 | 620 | 700 | 967 | 815
| Std Dev. |  | 7.5277 | 0 | 0 | 0 | 35.9258 | 10.4881
| Rel.Std.Dev. |  | 0.9549% | 0% | 0% | 0% | 3.772% | 1.287%
| PDM3 |  | -0.4715% | 11.1% | -21.72% | -11.62% | 20.23% | 2.895%
| Z-Score (Absolute) |  | 0.1 | 0.59 | 1.43 | 0.76 | 1.15 | 0.1
| Analytical results for Ti in OREAS 465 (Indicative Value 6.35 wt.%) | [8 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | wt.%
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 2
2
| 1 | 1 | 6.229 | 6.055 | 6.414 | 6.474 |  |  |  | 1
| 1 | 2 | 6.241 | 6.355 | 6.414 | 6.474 |  |  |  | 4
| 1 | 3 | 6.265 | 6.295 | 6.355 | 6.474 |  |  |  | 16
| 1 | 4 | 6.283 | 6.295 | 6.355 | 6.414 |  |  |  | 6.3505
| 1 | 5 | 6.265 | 6.235 | 6.414 | 6.474 |  |  |  | 10
| 1 | 6 | 6.271 | 6.295 | 6.414 | 6.414
| Mean |  | 6.259 | 6.255 | 6.3943 | 6.454
| Median |  | 6.265 | 6.295 | 6.414 | 6.474
| Std Dev. |  | 0.02008 | 0.10507 | 0.030467 | 0.030984
| Rel.Std.Dev. |  | 0.3208% | 1.68% | 0.4765% | 0.4801%
| PDM3 |  | -1.441% | -1.504% | 0.6896% | 1.629%
| Z-Score (Absolute) |  | 0.85 | 0.5 | 0.5 | 1.09
| Analytical results for Tm in OREAS 465 (Indicative Value < 10 ppm) | [8 empty] | Indicative
Tm | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 01 | 22 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 1
1
| 1 | 1 | < 90 | < 9 | [5 empty] | 1
| 1 | 2 | < 90 | < 9 | [5 empty] | 5
| 1 | 3 | < 90 | < 9 | [5 empty] | 16
| 1 | 4 | < 90 | < 9 | [5 empty] | < 10
| 1 | 5 | < 90 | < 9 | [5 empty] | 11
| 1 | 6 | < 90 | < 9
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for U in OREAS 465 (Indicative Value < 100 ppm) | [8 empty] | Indicative
U | Round | Replicate | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 09 | 13 | 14 | 15 |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF |  |  |  | 0
0
| 1 | 1 | < 80 | < 250 | < 80 | < 80 |  |  |  | 1
| 1 | 2 | < 80 | < 250 | < 80 | 80 |  |  |  | 6
| 1 | 3 | < 80 | < 250 | < 80 | < 80 |  |  |  | 16
| 1 | 4 | < 80 | < 250 | < 80 | 80 |  |  |  | < 100
| 1 | 5 | < 80 | < 250 | < 80 | 80 |  |  |  | 12
| 1 | 6 | < 80 | < 250 | < 80 | < 80
| Mean |  |  |  |  | 80
| Median |  |  |  |  | 80
| Std Dev. |  |  |  |  | 0
| Rel.Std.Dev. |  |  |  |  | 0%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA
| Analytical results for V in OREAS 465 (Indicative Value 609 ppm) | [8 empty] | Indicative
V | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 620 | 560 | [5 empty] | 1
| 1 | 2 | 670 | 560 | [5 empty] | 7
| 1 | 3 | 670 | 616 | [5 empty] | 16
| 1 | 4 | 670 | 532 | [5 empty] | 609.1754
| 1 | 5 | 670 | 560 | [5 empty] | 13
| 1 | 6 | 620 | 560
| Mean |  | 653.3333 | 564.6667
| Median |  | 670 | 560
| Std Dev. |  | 25.8199 | 27.5294
| Rel.Std.Dev. |  | 3.952% | 4.875%
| PDM3 |  | 7.249% | -7.306%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for W in OREAS 465 (Indicative Value < 100 ppm) | [8 empty] | Indicative
W | Round | Replicate | Lab | [6 empty] | 1
| No | No | 15 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | < 80 | [6 empty] | 1
| 1 | 2 | < 80 | [6 empty] | 8
| 1 | 3 | < 80 | [6 empty] | 16
| 1 | 4 | < 80 | [6 empty] | < 100
| 1 | 5 | < 80 | [6 empty] | 14
| 1 | 6 | < 80
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 465 (Certified Value 502 ppm) | [8 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 22 | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | BF*XRF | 0
0
| 1 | 1 | 400 | 550 | 470 | 550 | 550 | 500 | 465 | 1
| 1 | 2 | 400 | 550 | 550 | 470 | 470 | 500 | 488 | 0
| 1 | 3 | 400 | 550 | 470 | 470 | 470 | 500 | 472 | 16
| 1 | 4 | 400 | 470 | 470 | 470 | 550 | 500 | 472 | 502.0015
| 1 | 5 | 400 | 550 | 550 | 550 | 470 | 500 | 472 | 14
| 1 | 6 | 400 | 470 | 550 | 470 | 630 | 500 | 465
| Mean |  | 400 | 523.3333 | 510 | 496.6667 | 523.3333 | 500 | 472.3333
| Median |  | 400 | 550 | 510 | 470 | 510 | 500 | 472
| Std Dev. |  | 0 | 41.3118 | 43.8178 | 41.3118 | 65.3197 | 0 | 8.4063
| Rel.Std.Dev. |  | 0% | 7.894% | 8.592% | 8.318% | 12.48% | 0% | 1.78%
| PDM3 |  | -20.32% | 4.249% | 1.593% | -1.063% | 4.249% | -0.3987% | -5.91%
| Z-Score (Absolute) |  | 5.69 | 1.42 | 0.67 | 0.07 | 0.23 | 0 | 1.57
| Analytical results for Yb in OREAS 465 (Indicative Value 70 ppm) | [8 empty] | Indicative
Yb | Round | Replicate | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 14 | 22 |  |  |  |  | ppm
|  |  | BF*XRF | BF*XRF | BF*XRF |  |  |  |  | 0
0
| 1 | 1 | 90 | < 90 | 18 |  |  |  |  | 1
| 1 | 2 | 90 | 180 | 18 |  |  |  |  | 9
| 1 | 3 | < 90 | 90 | 18 |  |  |  |  | 16
| 1 | 4 | < 90 | 180 | 18 |  |  |  |  | 69.8249
| 1 | 5 | < 90 | 180 | 18 |  |  |  |  | 15
| 1 | 6 | 90 | 90 | 18
| Mean |  | 90 | 144 | 18
| Median |  | 90 | 180 | 18
| Std Dev. |  | 1.7405e-14 | 49.295 | 0
| Rel.Std.Dev. |  | 1.934e-14% | 34.23% | 0%
| PDM3 |  | 28.89% | 106.2% | -74.22%
| Z-Score (Absolute) |  | 0 | 0.77 | 0.67
| Analytical results for Zn in OREAS 465 (Indicative Value 1205 ppm) | [8 empty] | Indicative
Zn | Round | Replicate | Lab | [6 empty] | 1
| No | No | 14 | [6 empty] | ppm
|  |  | BF*XRF | [6 empty] | 0
0
| 1 | 1 | 1210 | [6 empty] | 1
| 1 | 2 | 1210 | [6 empty] | 10
| 1 | 3 | 1210 | [6 empty] | 16
| 1 | 4 | 1210 | [6 empty] | 1205.0129
| 1 | 5 | 1210 | [6 empty] | 16
| 1 | 6 | 1210
| Mean |  | 1210
| Median |  | 1210
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0.4139%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 465 (Indicative Value 1937 ppm) | [8 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | [5 empty] | 1
| No | No | 14 | 15 | [5 empty] | ppm
|  |  | BF*XRF | BF*XRF | [5 empty] | 0
0
| 1 | 1 | 2070 | 1920 | [5 empty] | 1
| 1 | 2 | 2000 | 1920 | [5 empty] | 11
| 1 | 3 | 1920 | 1920 | [5 empty] | 16
| 1 | 4 | 1920 | 1780 | [5 empty] | 1937.1237
| 1 | 5 | 2000 | 1920 | [5 empty] | 17
| 1 | 6 | 2000 | 1850
| Mean |  | 1985 | 1885
| Median |  | 2000 | 1920
| Std Dev. |  | 57.1839 | 58.5662
| Rel.Std.Dev. |  | 2.881% | 3.107%
| PDM3 |  | 2.472% | -2.691%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for H2O- in OREAS 465 (Indicative Value 0.312 wt.%) | [8 empty] | Indicative
H2O- | Round | Replicate | Lab | [6 empty] | 1
| No | No | 09 | [6 empty] | wt.%
|  |  | at 1000°C | [6 empty] | 3
3
| 1 | 1 | 0.31 | [6 empty] | 1
| 1 | 2 | 0.31 | [6 empty] | 13
| 1 | 3 | 0.3 | [6 empty] | 16
| 1 | 4 | 0.32 | [6 empty] | 0.31167
| 1 | 5 | 0.31 | [6 empty] | 19
| 1 | 6 | 0.32
| Mean |  | 0.31167
| Median |  | 0.31
| Std Dev. |  | 0.0075277
| Rel.Std.Dev. |  | 2.415%
| PDM3 |  | -8.882e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for LOI in OREAS 465 (Certified Value 0.824 wt.%) | [8 empty] | Certified
LOI | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 09 | 13 | 14 | 15 | 16 | 18 | wt.%
|  |  | at 1000°C | at 1000°C | at 1000°C | BF*XRF | at 1000°C | BF*XRF | at 1000°C | 3
3
| 1 | 1 | 0.78 | 0.68 | 0.64 | 0.94 | 1.08 | 0.752 | 0.94 | 1
| 1 | 2 | 0.8 | 0.68 | 0.68 | 0.83 | 1.08 | 0.711 | 0.91871 | 0
| 1 | 3 | 0.8 | 0.68 | 0.69 | 0.78 | 1.04 | 0.753 | 0.97843 | 16
| 1 | 4 | 0.76 | 0.68 | 0.75 | 0.85 | 1 | 0.759 | 0.91982 | 0.8238
| 1 | 5 | 0.74 | 0.71 | 0.72 | 0.9 | 1.07 | 0.774 | 0.97922 | 16
| 1 | 6 | 0.73 | 0.68 | 0.68 | 0.92 | 1.02 | 0.765 | 0.93981
| Mean |  | 0.76833 | 0.685 | 0.69333 | 0.87 | 1.0483 | 0.75233 | 0.946
| Median |  | 0.77 | 0.68 | 0.685 | 0.875 | 1.055 | 0.756 | 0.93991
| Std Dev. |  | 0.029944 | 0.012247 | 0.037771 | 0.060663 | 0.033714 | 0.021833 | 0.027055
| Rel.Std.Dev. |  | 3.897% | 1.788% | 5.448% | 6.973% | 3.216% | 2.902% | 2.86%
| PDM3 |  | -6.733% | -16.85% | -15.84% | 5.608% | 27.26% | -8.675% | 14.83%
| Z-Score (Absolute) |  | 0 | 0.67 | 0.57 | 0.78 | 2.14 | 0.06 | 1.36
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 465 (Indicative Value 6.5 ppm) | [16 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 20 | 22 | [12 empty] | ppm
|  |  | PF*MS | BF*MS | PF*MS | [12 empty] | 2
2
| 1 | 1 | 8 | 14 | 5 | [12 empty] | 1
| 1 | 2 | 8 | 47 | 5 | [12 empty] | 15
| 1 | 3 | 8 | 45 | 5 | [12 empty] | 16
| 1 | 4 | 8 | 13 | 5 | [12 empty] | 6.5
| 1 | 5 | 8 | 43 | 5 | [12 empty] | 21
| 1 | 6 | 8 | 40 | 5
| Mean |  | 8 | 33.6667 | 5
| Median |  | 8 | 41.5 | 5
| Std Dev. |  | 0 | 15.7945 | 0
| Rel.Std.Dev. |  | 0% | 46.91% | 0%
| PDM3 |  | 23.08% | 417.9% | -23.08%
| Z-Score (Absolute) |  | 0 | 5.77 | 0.67
| Analytical results for Al in OREAS 465 (Certified Value 6.6 wt.%) | [16 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 6.709 | 6.53 | 6.5 | 6.3616 | N.A. | 6.5746 | 5.96 | 6.87 | 7.03 | 6.85 | [5 empty] | 1
| 1 | 2 | 6.728 | 6.57 | 6.45 | 6.2861 | N.A. | 6.45 | 6.23 | 6.74 | 6.8 | 6.96 | [5 empty] | 1
| 1 | 3 | 6.715 | 6.5 | 6.48 | 6.3841 | N.A. | 6.7048 | 6.29 | 6.79 | 6.93 | 6.91 | [5 empty] | 16
| 1 | 4 | 6.607 | 6.57 | 6.55 | 6.3192 | N.A. | 6.6142 | 5.88 | 6.87 | 7 | 6.88 | [5 empty] | 6.5976
| 1 | 5 | 6.634 | 6.57 | 6.5 | 6.3793 | N.A. | 6.6482 | 6.2 | 6.83 | 6.87 | 6.87 | [5 empty] | 18
| 1 | 6 | 6.607 | 6.5 | 6.62 | 6.271 | -- | 6.5689 | 6.1 | 6.75 | 6.89 | 6.87
| Mean |  | 6.6667 | 6.54 | 6.5167 | 6.3335 |  | 6.5935 | 6.11 | 6.8083 | 6.92 | 6.89
| Median |  | 6.6715 | 6.55 | 6.5 | 6.3404 |  | 6.5944 | 6.15 | 6.81 | 6.91 | 6.875
| Std Dev. |  | 0.056705 | 0.034641 | 0.060222 | 0.048559 |  | 0.086453 | 0.16149 | 0.057417 | 0.085323 | 0.039497
| Rel.Std.Dev. |  | 0.8506% | 0.5297% | 0.9241% | 0.7667% |  | 1.311% | 2.643% | 0.8433% | 1.233% | 0.5732%
| PDM3 |  | 1.046% | -0.8735% | -1.227% | -4.003% |  | -0.06303% | -7.391% | 3.194% | 4.886% | 4.431%
| Z-Score (Absolute) |  | 0.23 | 0.17 | 0.24 | 0.82 | NA | 0 | 1.52 | 0.67 | 1.02 | 0.93
| Analytical results for As in OREAS 465 (Indicative Value 420 ppm) | [16 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | BF*OES/MS | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | 31 | < 100 | 354.1107 | 570 | 70 | 1070 | [9 empty] | 1
| 1 | 2 | 39 | < 100 | 362.3183 | 570 | 70 | 1080 | [9 empty] | 16
| 1 | 3 | 24 | < 100 | 350.9054 | 590 | 70 | 1070 | [9 empty] | 16
| 1 | 4 | 35 | < 100 | 384.4084 | 570 | 70 | 1020 | [9 empty] | 419.6712
| 1 | 5 | 36 | < 100 | 373.3242 | 560 | 70 | 1070 | [9 empty] | 22
| 1 | 6 | 32 | < 100 | 358.0684 | 570 | 60 | 1050
| Mean |  | 32.8333 |  | 363.8559 | 571.6667 | 68.3333 | 1060
| Median |  | 33.5 |  | 360.1934 | 570 | 70 | 1070
| Std Dev. |  | 5.1929 |  | 12.7326 | 9.8319 | 4.0825 | 21.9089
| Rel.Std.Dev. |  | 15.82% |  | 3.499% | 1.72% | 5.974% | 2.067%
| PDM3 |  | -92.18% |  | -13.3% | 36.22% | -83.72% | 152.6%
| Z-Score (Absolute) |  | 0.76 | NA | 0 | 0.48 | 0.67 | 1.6
| Analytical results for B in OREAS 465 (Indicative Value 23.3 ppm) | [16 empty] | Indicative
B | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*OES | PF*OES | [13 empty] | 1
1
| 1 | 1 | < 50 | 20 | [13 empty] | 1
| 1 | 2 | < 50 | 40 | [13 empty] | 17
| 1 | 3 | < 50 | 20 | [13 empty] | 16
| 1 | 4 | < 50 | 20 | [13 empty] | 23.3333
| 1 | 5 | < 50 | 20 | [13 empty] | 23
| 1 | 6 | < 50 | 20
| Mean |  |  | 23.3333
| Median |  |  | 20
| Std Dev. |  |  | 8.165
| Rel.Std.Dev. |  |  | 34.99%
| PDM3 |  |  | 1.332e-13%
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Ba in OREAS 465 (Certified Value 4397 ppm) | [16 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*OES | BF*OES | PF*OES | BF*OES/MS | BF*OES | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 4760 | 4590 | 4650 | 4640 | 4840 | 4265.1643 | 4304 | 4296.5 | 3685.424 | N.A. | 4650 | 4097.8907 | 4350 | 5400 | 3530 | 1
| 1 | 2 | 4610 | 3990 | 4670 | 4750 | 4680 | 4208.7488 | 4242 | 4305.7 | 3612.24 | N.A. | 4550 | 4275.1938 | 4210 | 5150 | 3530 | 2
| 1 | 3 | 4870 | 4120 | 4720 | 4740 | 4890 | 4283.7736 | 4222 | 4346.8 | 3792.032 | N.A. | 4600 | 4271.3152 | 4220 | 5230 | 3620 | 16
| 1 | 4 | 4820 | 4230 | 4600 | 4780 | 4790 | 4159.7898 | 4184 | 4333.6 | 3665.528 | N.A. | 4650 | 4309.8525 | 4240 | 5770 | 3520 | 4397.0592
| 1 | 5 | 4760 | 4060 | 4650 | 4780 | 4640 | 4251.8818 | 4187 | 4278.9 | 3712.112 | N.A. | 4600 | 4257.351 | 4200 | 5250 | 3540 | 19
| 1 | 6 | 4820 | 4090 | 4670 | 4650 | 4740 | 4260.7857 | 4250 | 4426 | 3675.488 | -- | 4550 | 4365.8776 | 4230 | 5410 | 3520
| Mean |  | 4773.3333 | 4180 | 4660 | 4723.3333 | 4763.3333 | 4238.3573 | 4231.5 | 4331.25 | 3690.4707 |  | 4600 | 4262.9135 | 4241.6667 | 5368.3333 | 3543.3333
| Median |  | 4790 | 4105 | 4660 | 4745 | 4765 | 4256.3337 | 4232 | 4319.65 | 3680.456 |  | 4600 | 4273.2545 | 4225 | 5325 | 3530
| Std Dev. |  | 90.2589 | 215.7777 | 38.9872 | 62.8225 | 95.219 | 45.8622 | 44.8007 | 52.5847 | 59.6241 |  | 44.7214 | 89.7661 | 54.9242 | 221.3067 | 38.2971
| Rel.Std.Dev. |  | 1.891% | 5.162% | 0.8366% | 1.33% | 1.999% | 1.082% | 1.059% | 1.214% | 1.616% |  | 0.9722% | 2.106% | 1.295% | 4.122% | 1.081%
| PDM3 |  | 8.557% | -4.936% | 5.98% | 7.42% | 8.33% | -3.609% | -3.765% | -1.497% | -16.07% |  | 4.615% | -3.051% | -3.534% | 22.09% | -19.42%
| Z-Score (Absolute) |  | 0.98 | 0.46 | 0.74 | 0.87 | 0.96 | 0.16 | 0.17 | 0.04 | 1.33 | NA | 0.61 | 0.04 | 0.15 | 2.25 | 1.64
| Analytical results for Be in OREAS 465 (Certified Value 13.2 ppm) | [16 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 17 | 22 | 23 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | [8 empty] | 1
2
| 1 | 1 | 13.297 | 14 | 13.3 | 9.0083 | 13 | 14 | 14 | [8 empty] | 1
| 1 | 2 | 13.4222 | 14 | 12.8 | 8.8068 | 14 | 14 | 12 | [8 empty] | 0
| 1 | 3 | 13.2681 | 13 | 12.2 | 8.8283 | 14 | 13 | 11 | [8 empty] | 16
| 1 | 4 | 13.9602 | 14 | 13.1 | 9.0659 | 13 | 15 | 11 | [8 empty] | 13.246
| 1 | 5 | 13.0956 | 14 | 13.8 | 8.5694 | 14 | 13 | 10 | [8 empty] | 20
| 1 | 6 | 13.4144 | 14 | 13.2 | 8.8828 | 12 | 14 | 12
| Mean |  | 13.4096 | 13.8333 | 13.0667 | 8.8603 | 13.3333 | 13.8333 | 11.6667
| Median |  | 13.3557 | 14 | 13.15 | 8.8556 | 13.5 | 14 | 11.5
| Std Dev. |  | 0.29483 | 0.40825 | 0.53541 | 0.17504 | 0.8165 | 0.75277 | 1.3663
| Rel.Std.Dev. |  | 2.199% | 2.951% | 4.098% | 1.976% | 6.124% | 5.442% | 11.71%
| PDM3 |  | 1.235% | 4.434% | -1.354% | -33.11% | 0.659% | 4.434% | -11.92%
| Z-Score (Absolute) |  | 0.1 | 0.67 | 0.36 | 6.03 | 0 | 0.67 | 1.8
| Analytical results for Bi in OREAS 465 (Certified Value 16.9 ppm) | [16 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 17.4119 | 17.5 | 17 | 8.3 | 17 | 16.1 | [9 empty] | 1
| 1 | 2 | 17.271 | 17.6 | 16.5 | 10.5 | 15 | 15.6 | [9 empty] | 0
| 1 | 3 | 17.5399 | 17.1 | 16.5 | 10.6 | 17 | 15.7 | [9 empty] | 16
| 1 | 4 | 17.8256 | 17.3 | 17 | 8 | 20 | 15.8 | [9 empty] | 16.9218
| 1 | 5 | 16.8947 | 17.4 | 17 | 10.7 | 17 | 15.6 | [9 empty] | 21
| 1 | 6 | 17.5102 | 17.2 | 16 | 10.4 | 19 | 15.3
| Mean |  | 17.4089 | 17.35 | 16.6667 | 9.75 | 17.5 | 15.6833
| Median |  | 17.4611 | 17.35 | 16.75 | 10.45 | 17 | 15.65
| Std Dev. |  | 0.31128 | 0.18708 | 0.40825 | 1.247 | 1.7607 | 0.26394
| Rel.Std.Dev. |  | 1.788% | 1.078% | 2.449% | 12.79% | 10.06% | 1.683%
| PDM3 |  | 2.879% | 2.531% | -1.508% | -42.38% | 3.417% | -7.319%
| Z-Score (Absolute) |  | 0.61 | 0.52 | 0.52 | 10.97 | 0.74 | 2
| Analytical results for Ca in OREAS 465 (Certified Value 0.9 wt.%) | [16 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.895 | 0.9 | 0.98 | 0.79039 | N.A. | 1.02 | 0.87134 | 0.89 | 0.966 | 1 | 0.8 |  |  |  |  | 1
| 1 | 2 | 0.915 | 0.9 | 0.94 | 0.79575 | N.A. | 0.9 | 0.87491 | 0.89 | 0.926 | 0.9 | 0.9 |  |  |  |  | 3
| 1 | 3 | 0.912 | 0.9 | 0.96 | 0.81722 | N.A. | 1.01 | 0.89064 | 0.9 | 0.96 | 1 | 0.8 |  |  |  |  | 16
| 1 | 4 | 0.879 | 0.9 | 0.94 | 0.79403 | N.A. | 0.87 | 0.87563 | 0.9 | 0.946 | 1 | 0.8 |  |  |  |  | 0.89974
| 1 | 5 | 0.878 | 0.9 | 0.95 | 0.8092 | N.A. | 0.96 | 0.88635 | 0.9 | 0.955 | 0.9 | 0.8 |  |  |  |  | 22
| 1 | 6 | 0.812 | 0.9 | 0.97 | 0.79657 | -- | 0.96 | 0.86919 | 0.9 | 0.929 | 0.9 | 0.9
| Mean |  | 0.88183 | 0.9 | 0.95667 | 0.80053 |  | 0.95333 | 0.87801 | 0.89667 | 0.947 | 0.95 | 0.83333
| Median |  | 0.887 | 0.9 | 0.955 | 0.79616 |  | 0.96 | 0.87527 | 0.9 | 0.9505 | 0.95 | 0.8
| Std Dev. |  | 0.037637 | 1.2162e-16 | 0.01633 | 0.010369 |  | 0.059217 | 0.0085617 | 0.005164 | 0.016492 | 0.054772 | 0.05164
| Rel.Std.Dev. |  | 4.268% | 1.351e-14% | 1.707% | 1.295% |  | 6.212% | 0.9751% | 0.5759% | 1.742% | 5.766% | 6.197%
| PDM3 |  | -1.99% | 0.02924% | 6.327% | -11.03% |  | 5.957% | -2.415% | -0.3412% | 5.253% | 5.586% | -7.38%
| Z-Score (Absolute) |  | 0.22 | 0.02 | 0.78 | 1.31 | NA | 0.74 | 0.27 | 0.02 | 0.65 | 0.69 | 0.87
| Analytical results for Cd in OREAS 465 (Indicative Value 1.04 ppm) | [16 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | PF*OES/MS | [11 empty] | 2
2
| 1 | 1 | 1 | 1 | < 10 | 1.3 | [11 empty] | 1
| 1 | 2 | 1 | < 1 | < 10 | 1.2 | [11 empty] | 18
| 1 | 3 | 1 | < 1 | < 10 | 1.4 | [11 empty] | 16
| 1 | 4 | 2 | 1 | < 10 | 1.1 | [11 empty] | 1.0389
| 1 | 5 | 1 | 1 | < 10 | 1.5 | [11 empty] | 24
| 1 | 6 | 1 | 1 | < 10 | 1.2
| Mean |  | 1.1667 | 1 |  | 1.2833
| Median |  | 1 | 1 |  | 1.25
| Std Dev. |  | 0.40825 | 0 |  | 0.1472
| Rel.Std.Dev. |  | 34.99% | 0% |  | 11.47%
| PDM3 |  | 12.3% | -3.743% |  | 23.53%
| Z-Score (Absolute) |  | 0 | 0.67 | NA | 1.15
| Analytical results for Ce in OREAS 465 (Certified Value 3.95 wt.%) | [16 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*OES | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.22 | 3.89 | 3.84 | > 1 | 3.9 | 3.6629 | 3.9527 | 3.8853 | 4.0546 | 3.851 | 4.0385 | 4.0949 | 1.44 | 4 | 4.34 | 1
| 1 | 2 | 4.31 | 3.85 | 3.85 | > 1 | 3.89 | 3.6604 | 3.9988 | 3.8573 | 3.925 | 3.81 | 4.0246 | 4.2829 | 4.04 | 3.95 | 4.31 | 0
| 1 | 3 | 4.28 | 4.05 | 3.84 | > 1 | 3.99 | 3.7241 | 3.9294 | 3.8762 | 3.9377 | 3.818 | 4.0376 | 4.2883 | 4.06 | 3.94 | 4.34 | 16
| 1 | 4 | 4.09 | 4.03 | 3.94 | > 1 | 3.83 | 3.7184 | 3.9881 | 3.9105 | 3.9792 | 3.826 | 4.043 | 4.3143 | 4.16 | 4 | 4.26 | 3.9525
| 1 | 5 | 4.26 | 3.88 | 3.93 | > 1 | 3.9 | 3.717 | 3.967 | 3.8797 | 3.9621 | 3.826 | 4.2349 | 4.2852 | 4.14 | 3.95 | 4.28 | 23
| 1 | 6 | 4.09 | 4.06 | 3.88 | > 1 | 4.01 | 3.7308 | 3.9317 | 3.9396 | 4.0055 | 3.794 | 3.9694 | 4.4003 | 4.13 | 4.04 | 4.28
| Mean |  | 4.2083 | 3.96 | 3.88 |  | 3.92 | 3.7023 | 3.9613 | 3.8914 | 3.9773 | 3.8208 | 4.058 | 4.2776 | 3.6617 | 3.98 | 4.3017
| Median |  | 4.24 | 3.96 | 3.865 |  | 3.9 | 3.7177 | 3.9598 | 3.8825 | 3.9707 | 3.822 | 4.0381 | 4.2867 | 4.095 | 3.975 | 4.295
| Std Dev. |  | 0.096212 | 0.096333 | 0.045166 |  | 0.067528 | 0.031857 | 0.028722 | 0.029175 | 0.04756 | 0.01902 | 0.090863 | 0.099984 | 1.0894 | 0.039497 | 0.033714
| Rel.Std.Dev. |  | 2.286% | 2.433% | 1.164% |  | 1.723% | 0.8605% | 0.7251% | 0.7497% | 1.196% | 0.4978% | 2.239% | 2.337% | 29.75% | 0.9924% | 0.7838%
| PDM3 |  | 6.473% | 0.1898% | -1.834% |  | -0.8222% | -6.331% | 0.2223% | -1.545% | 0.6285% | -3.331% | 2.669% | 8.226% | -7.358% | 0.6958% | 8.834%
| Z-Score (Absolute) |  | 1.93 | 0.08 | 0.72 | NA | 0.4 | 2.15 | 0.06 | 0.63 | 0.06 | 1.2 | 0.43 | 2.78 | 1.1 | 0.09 | 2.68
| Analytical results for Co in OREAS 465 (Indicative Value 23 ppm) | [16 empty] | Indicative
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 1
1
| 1 | 1 | 24.5 | 27.0248 | N.A. | 20.1304 | 9.1 | < 10 | < 20 | 18.9 | [7 empty] | 1
| 1 | 2 | 23.9 | 27.3713 | N.A. | 23.224 | 14.4 | < 10 | 20 | 18.9 | [7 empty] | 19
| 1 | 3 | 24.2 | 27.1743 | N.A. | 21.0304 | 116 | < 10 | < 20 | 18.8 | [7 empty] | 16
| 1 | 4 | 24 | 27.7312 | N.A. | 22.8897 | 50.5 | < 10 | < 20 | 17.4 | [7 empty] | 22.9817
| 1 | 5 | 25 | 26.9341 | N.A. | 22.3463 | 31.8 | < 10 | < 20 | 18.5 | [7 empty] | 25
| 1 | 6 | 24.5 | 26.5791 | -- | 33.7247 | 48.7 | < 10 | 20 | 18.6
| Mean |  | 24.35 | 27.1358 |  | 23.8909 | 45.0833 |  | 20 | 18.5167
| Median |  | 24.35 | 27.0996 |  | 22.618 | 40.25 |  | 20 | 18.7
| Std Dev. |  | 0.40373 | 0.39352 |  | 4.957 | 38.6893 |  | 0 | 0.57067
| Rel.Std.Dev. |  | 1.658% | 1.45% |  | 20.75% | 85.82% |  | 0% | 3.082%
| PDM3 |  | 5.954% | 18.08% |  | 3.956% | 96.17% |  | -12.97% | -19.43%
| Z-Score (Absolute) |  | 0.19 | 0.63 | NA | 0.19 | 1.21 | NA | 1.53 | 0.72
| Analytical results for Cr in OREAS 465 (Certified Value 544 ppm) | [16 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | 0
0
| 1 | 1 | 550 | 560 | 600 | 610 | 600 | < 1000 | 536 | 526.503 | N.A. | 550 | 485.1 | < 100 | 1260 | 260 | 560 | 1
| 1 | 2 | 560 | 480 | 600 | 620 | 580 | < 1000 | 492 | 528.967 | N.A. | 600 | 485.8 | 500 | 1180 | 260 | 560 | 4
| 1 | 3 | 540 | 490 | 580 | 620 | 600 | < 1000 | 539 | 525.483 | N.A. | 550 | 479.6 | 500 | 1660 | 280 | 550 | 16
| 1 | 4 | 540 | 510 | 560 | 650 | 560 | < 1000 | 530 | 528.471 | N.A. | 550 | 499.5 | < 100 | 1230 | 280 | 560 | 544.1822
| 1 | 5 | 530 | 500 | 580 | 650 | 580 | < 1000 | 557 | 528.155 | N.A. | 550 | 489.9 | 500 | 1200 | 270 | 560 | 24
| 1 | 6 | 520 | 500 | 560 | 590 | 580 | < 1000 | 540 | 529.264 | -- | 550 | 489.2 | 500 | 1180 | 290 | 550
| Mean |  | 540 | 506.6667 | 580 | 623.3333 | 583.3333 |  | 532.3333 | 527.8072 |  | 558.3333 | 488.1833 | 500 | 1285 | 273.3333 | 556.6667
| Median |  | 540 | 500 | 580 | 620 | 580 |  | 537.5 | 528.313 |  | 550 | 487.5 | 500 | 1215 | 275 | 560
| Std Dev. |  | 14.1421 | 28.0476 | 17.8885 | 23.3809 | 15.0555 |  | 21.7133 | 1.4921 |  | 20.4124 | 6.6469 | 0 | 186.3062 | 12.1106 | 5.164
| Rel.Std.Dev. |  | 2.619% | 5.536% | 3.084% | 3.751% | 2.581% |  | 4.079% | 0.2827% |  | 3.656% | 1.362% | 0% | 14.5% | 4.431% | 0.9277%
| PDM3 |  | -0.7685% | -6.894% | 6.582% | 14.54% | 7.194% |  | -2.177% | -3.009% |  | 2.6% | -10.29% | -8.119% | 136.1% | -49.77% | 2.294%
| Z-Score (Absolute) |  | 0 | 0.74 | 0.67 | 1.4 | 0.73 | NA | 0.13 | 0.21 | NA | 0.31 | 0.87 | 0.67 | 11.29 | 4.5 | 0.28
| Analytical results for Cs in OREAS 465 (Certified Value < 0.1 ppm) | [16 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
2
| 1 | 1 | 0.09 | 0.02 | 0.07 | < 0.01 | 0.15 | < 0.1 | < 0.1 | 0.3 | N.A. | < 1 | < 1 | 0.3 | < 1 | 0.2 |  | 1
| 1 | 2 | 0.1 | < 0.01 | 0.07 | 0.01 | 0.1 | < 0.1 | < 0.1 | 0.2 | N.A. | < 1 | < 1 | 0.1 | < 1 | 0.3 |  | 0
| 1 | 3 | 0.16 | 0.05 | 0.12 | < 0.01 | 0.09 | < 0.1 | < 0.1 | 0.3 | N.A. | < 1 | < 1 | < 0.1 | < 1 | 0.4 |  | 16
| 1 | 4 | 0.08 | 0.01 | 0.11 | < 0.01 | < 0.01 | < 0.1 | < 0.1 | 0.3 | N.A. | < 1 | < 1 | < 0.1 | < 1 | 0.3 |  | < 0.1
| 1 | 5 | 0.06 | 0.04 | 0.1 | < 0.01 | < 0.01 | < 0.1 | < 0.1 | 0.3 | N.A. | < 1 | < 1 | < 0.1 | < 1 | 0.2 |  | 25
| 1 | 6 | 0.07 | 0.02 | 0.08 | < 0.01 | 0.06 | < 0.1 | < 0.1 | 0.3 | -- | < 1 | < 1 | < 0.1 | < 1 | 0.2
| Mean |  | 0.093333 | 0.028 | 0.091667 | 0.01 | 0.1 |  |  | 0.28333 |  |  |  | 0.2 |  | 0.26667
| Median |  | 0.085 | 0.02 | 0.09 | 0.01 | 0.095 |  |  | 0.3 |  |  |  | 0.2 |  | 0.25
| Std Dev. |  | 0.03559 | 0.016432 | 0.02137 |  | 0.037417 |  |  | 0.040825 |  |  |  | 0.14142 |  | 0.08165
| Rel.Std.Dev. |  | 38.13% | 58.68% | 23.31% |  | 37.42% |  |  | 14.41% |  |  |  | 70.71% |  | 30.62%
PDM3
| Z-Score (Absolute) |  | 0.21 | 0.97 | 0.21 | 1.36 | 0.38 | NA | NA | 4.51 | NA | NA | NA | 0.21 | NA | 4.16
| Analytical results for Cu in OREAS 465 (Indicative Value 128 ppm) | [16 empty] | Indicative
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [8 empty] | ppm
|  |  | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [8 empty] | 0
0
| 1 | 1 | N.A. | 100 | 165.7163 | 43 | 120 | 100 | 170 | [8 empty] | 1
| 1 | 2 | N.A. | 100 | 197.9965 | 218 | 110 | 100 | 170 | [8 empty] | 20
| 1 | 3 | N.A. | 100 | 182.8171 | 97 | 110 | 100 | 160 | [8 empty] | 16
| 1 | 4 | N.A. | 200 | 179.7828 | 50 | 110 | 100 | 170 | [8 empty] | 128.3315
| 1 | 5 | N.A. | 200 | 182.0189 | 76 | 120 | 100 | 170 | [8 empty] | 26
| 1 | 6 | -- | 100 | 207.6015 | 79 | 110 | 100 | 170
| Mean |  |  | 133.3333 | 185.9888 | 93.8333 | 113.3333 | 100 | 168.3333
| Median |  |  | 100 | 182.418 | 77.5 | 110 | 100 | 170
| Std Dev. |  |  | 51.6398 | 14.7418 | 63.9857 | 5.164 | 0 | 4.0825
| Rel.Std.Dev. |  |  | 38.73% | 7.926% | 68.19% | 4.556% | 0% | 2.425%
| PDM3 |  |  | 3.898% | 44.93% | -26.88% | -11.69% | -22.08% | 31.17%
| Z-Score (Absolute) |  | NA | 0.2 | 1.24 | 1.07 | 0.2 | 0.46 | 0.89
| Analytical results for Dy in OREAS 465 (Certified Value 217 ppm) | [16 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 213 | 214 | 218 | 219 | 212 | 223.4385 | 218.5 | 216.8 | 179.0049 | 170 | 223 | 192.6 | 60.2 | 226 | 210 | 1
| 1 | 2 | 219 | 212 | 230 | 221 | 217 | 226.0935 | 216.3 | 211.7 | 180.1796 | 170 | 222.5 | 186.5 | 183 | 219 | 202 | 5
| 1 | 3 | 218 | 214 | 223 | 225 | 215 | 233.2634 | 215 | 210.3 | 179.3258 | 170 | 220 | 196 | 20.8 | 228 | 209 | 16
| 1 | 4 | 209 | 224 | 243 | 225 | 208 | 224.2843 | 216.8 | 216.8 | 180.4077 | 170 | 218.5 | 189.9 | 27.8 | 252 | 204 | 216.7238
| 1 | 5 | 217 | 211 | 239 | 223 | 218 | 225.1254 | 213.8 | 216.3 | 180.5249 | 170 | 233 | 191.7 | 93.1 | 219 | 202 | 26
| 1 | 6 | 205 | 209 | 245 | 220 | 219 | 231.1552 | 215.2 | 217.9 | 177.6372 | 170 | 217 | 186.5 | 98.4 | 242 | 205
| Mean |  | 213.5 | 214 | 233 | 222.1667 | 214.8333 | 227.2267 | 215.9333 | 214.9667 | 179.5134 | 170 | 222.3333 | 190.5333 | 80.55 | 231 | 205.3333
| Median |  | 215 | 213 | 234.5 | 222 | 216 | 225.6095 | 215.75 | 216.55 | 179.7527 | 170 | 221.25 | 190.8 | 76.65 | 227 | 204.5
| Std Dev. |  | 5.5767 | 5.2536 | 11.0815 | 2.5626 | 4.1673 | 4.0146 | 1.6391 | 3.1481 | 1.1028 | 0 | 5.7067 | 3.7001 | 59.5632 | 13.2966 | 3.4448
| Rel.Std.Dev. |  | 2.612% | 2.455% | 4.756% | 1.153% | 1.94% | 1.767% | 0.7591% | 1.464% | 0.6143% | 0% | 2.567% | 1.942% | 73.95% | 5.756% | 1.678%
| PDM3 |  | -1.488% | -1.257% | 7.51% | 2.511% | -0.8723% | 4.846% | -0.3648% | -0.8108% | -17.17% | -21.56% | 2.588% | -12.08% | -62.83% | 6.587% | -5.256%
| Z-Score (Absolute) |  | 0.09 | 0.2 | 1.29 | 0.52 | 0 | 0.88 | 0.08 | 0.01 | 2.51 | 2.88 | 0.38 | 1.73 | 9.53 | 1.15 | 0.67
| Analytical results for Er in OREAS 465 (Certified Value 50 ppm) | [16 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 46.9 | 52.8 | 53.8 | 51.8 | 50.9 | 50.3491 | 49.2 | 48.9 | 56.2029 | 44 | 49.1 | 40.1008 | 7.19 | 51.5 | 52.6 | 1
| 1 | 2 | 46.1 | 48.5 | 54.9 | 52.6 | 47.6 | 50.7643 | 50.3 | 48.7 | 56.9099 | 44 | 49.4 | 42.0845 | 9.82 | 51 | 52.1 | 6
| 1 | 3 | 49.2 | 52.2 | 52.5 | 51.9 | 50.8 | 50.9489 | 48.2 | 47.7 | 54.6547 | 44 | 49.5 | 41.7552 | 4.75 | 51.5 | 51.4 | 16
| 1 | 4 | 49.7 | 48.2 | 56.4 | 52.1 | 48.8 | 48.6333 | 49.7 | 49.1 | 54.5386 | 44 | 48.6 | 41.362 | 8.13 | 59.5 | 54.7 | 50.459
| 1 | 5 | 49.7 | 51.5 | 53.7 | 51.6 | 49.2 | 50.8571 | 49.4 | 48.4 | 57.9494 | 44 | 51 | 41.5998 | 11.2 | 49.5 | 52.1 | 27
| 1 | 6 | 50.4 | 49.6 | 53.5 | 50.7 | 48.8 | 50.5171 | 49.3 | 48.8 | 57.4205 | 44 | 48.8 | 42.5537 | 11.8 | 54 | 51.3
| Mean |  | 48.6667 | 50.4667 | 54.1333 | 51.7833 | 49.35 | 50.345 | 49.35 | 48.6 | 56.2794 | 44 | 49.4 | 41.576 | 8.815 | 52.8333 | 52.3667
| Median |  | 49.45 | 50.55 | 53.75 | 51.85 | 49 | 50.6407 | 49.35 | 48.75 | 56.5564 | 44 | 49.25 | 41.6775 | 8.975 | 51.5 | 52.1
| Std Dev. |  | 1.7397 | 1.9633 | 1.3486 | 0.63061 | 1.2802 | 0.86746 | 0.6892 | 0.498 | 1.4258 | 0 | 0.85557 | 0.83297 | 2.6549 | 3.573 | 1.242
| Rel.Std.Dev. |  | 3.575% | 3.89% | 2.491% | 1.218% | 2.594% | 1.723% | 1.397% | 1.025% | 2.533% | 0% | 1.732% | 2.003% | 30.12% | 6.763% | 2.372%
| PDM3 |  | -3.552% | 0.01529% | 7.282% | 2.625% | -2.198% | -0.2259% | -2.198% | -3.684% | 11.53% | -12.8% | -2.099% | -17.6% | -82.53% | 4.706% | 3.781%
| Z-Score (Absolute) |  | 0.24 | 0.34 | 1.52 | 0.77 | 0.02 | 0.3 | 0.02 | 0.26 | 2.21 | 1.82 | 0 | 2.51 | 13.03 | 0.67 | 0.95
| Analytical results for Eu in OREAS 465 (Certified Value 286 ppm) | [16 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 288 | 294 | 290 | 279 | 294 | 272.4468 | 298.6 | 275.7 | 181.6945 | 285 | 283 | 270.9363 | 91.3 | 292 | 275 | 1
| 1 | 2 | 294 | 285 | 292 | 284 | 294 | 275.2665 | 299.4 | 276.3 | 183.4952 | 276 | 281 | 284.3672 | 336 | 277 | 272 | 7
| 1 | 3 | 295 | 301 | 292 | 288 | 296 | 278.1256 | 295 | 267.2 | 182.3763 | 276 | 281 | 284.8866 | 76.9 | 290 | 278 | 16
| 1 | 4 | 277 | 290 | 330 | 287 | 283 | 271.7197 | 293.8 | 272.4 | 183.9973 | 276 | 277 | 286.1175 | 140 | 317 | 273 | 285.6091
| 1 | 5 | 293 | 295 | 321 | 284 | 295 | 276.9184 | 294 | 273.7 | 183.8519 | 276 | 295.5 | 284.8958 | 76.7 | 281 | 268 | 28
| 1 | 6 | 281 | 295 | 321 | 279 | 299 | 276.6569 | 295.6 | 273 | 182.1157 | 276 | 275 | 291.0592 | 98 | 304 | 270
| Mean |  | 288 | 293.3333 | 307.6667 | 283.5 | 293.5 | 275.189 | 296.0667 | 273.05 | 182.9218 | 277.5 | 282.0833 | 283.7105 | 136.4833 | 293.5 | 272.6667
| Median |  | 290.5 | 294.5 | 306.5 | 284 | 294.5 | 275.9617 | 295.3 | 273.35 | 182.9358 | 276 | 281 | 284.8912 | 94.65 | 291 | 272.5
| Std Dev. |  | 7.4833 | 5.3914 | 18.2062 | 3.8341 | 5.4681 | 2.5819 | 2.3788 | 3.2452 | 0.98023 | 3.6742 | 7.2001 | 6.726 | 100.4649 | 14.8694 | 3.559
| Rel.Std.Dev. |  | 2.598% | 1.838% | 5.918% | 1.352% | 1.863% | 0.9382% | 0.8035% | 1.188% | 0.5359% | 1.324% | 2.552% | 2.371% | 73.61% | 5.066% | 1.305%
| PDM3 |  | 0.8371% | 2.704% | 7.723% | -0.7384% | 2.763% | -3.648% | 3.662% | -4.397% | -35.95% | -2.839% | -1.234% | -0.6648% | -52.21% | 2.763% | -4.532%
| Z-Score (Absolute) |  | 0.3 | 0.66 | 1.33 | 0 | 0.67 | 0.56 | 0.85 | 0.7 | 6.78 | 0.38 | 0.28 | 0.01 | 12.6 | 0.67 | 0.73
| Analytical results for Fe in OREAS 465 (Certified Value 34.71 wt.%) | [16 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 32.3318 | 34.11 | 34.37 | 34.4032 | N.A. | 35.56 | 34.1539 | 8.743 | 35.1 | 35.9 | > 25 |  |  |  |  | 1
| 1 | 2 | 32.1486 | 34.58 | 34.09 | 34.2612 | N.A. | 35.17 | 33.8011 | 30.986 | 34.71 | 35 | > 25 |  |  |  |  | 0
| 1 | 3 | 32.6404 | 33.54 | 34.21 | 34.3024 | N.A. | 35.48 | 34.4999 | 31.195 | 34.56 | 35.8 | > 25 |  |  |  |  | 16
| 1 | 4 | 31.4639 | 34.25 | 34.73 | 34.2361 | N.A. | 35.95 | 34.3167 | 7.554 | 34.35 | 35.8 | > 25 |  |  |  |  | 34.7122
| 1 | 5 | 31.5989 | 34.16 | 34.42 | 34.6676 | N.A. | 35.59 | 34.4117 | 30.986 | 34.61 | 35.4 | > 25 |  |  |  |  | 29
| 1 | 6 | 31.6182 | 33.9 | 34.87 | 35.2713 | -- | 35.78 | 33.9978 | 29.517 | 34.6 | 35 | > 25
| Mean |  | 31.967 | 34.09 | 34.4483 | 34.5236 |  | 35.5883 | 34.1968 | 23.1635 | 34.655 | 35.4833
| Median |  | 31.8834 | 34.135 | 34.395 | 34.3528 |  | 35.575 | 34.2353 | 30.2515 | 34.605 | 35.6
| Std Dev. |  | 0.47532 | 0.34917 | 0.29976 | 0.39868 |  | 0.26649 | 0.26475 | 11.6521 | 0.24825 | 0.41191
| Rel.Std.Dev. |  | 1.487% | 1.024% | 0.8702% | 1.155% |  | 0.7488% | 0.7742% | 50.3% | 0.7164% | 1.161%
| PDM3 |  | -7.909% | -1.792% | -0.7602% | -0.5433% |  | 2.524% | -1.485% | -33.27% | -0.1648% | 2.221%
| Z-Score (Absolute) |  | 4.67 | 0.67 | 0 | 0.14 | NA | 2.15 | 0.47 | 21.24 | 0.39 | 1.95 | NA
| Analytical results for Ga in OREAS 465 (Certified Value 214 ppm) | [16 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 05 | 06 | 08 | 12 | 14 | 20 | 22 | 23 | [5 empty] | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*MS | PF*OES/MS | [5 empty] | 0
0
| 1 | 1 | 200 | 204 | 222 | 388 | 589.2254 | 143.5778 | N.A. | < 1 | 290 | 200 | [5 empty] | 1
| 1 | 2 | 225 | 204 | 223 | 373 | 590.498 | 141.8727 | N.A. | < 1 | 274 | 198 | [5 empty] | 0
| 1 | 3 | 235 | 204 | 226 | 392 | 588.9666 | 142.9924 | N.A. | < 1 | 278 | 202 | [5 empty] | 16
| 1 | 4 | 240 | 198.5 | 228 | 353 | 588.2637 | 141.6857 | N.A. | < 1 | 290 | 198 | [5 empty] | 213.7519
| 1 | 5 | 233 | 198.5 | 228 | 364 | 591.7189 | 141.4113 | N.A. | < 1 | 280 | 208 | [5 empty] | 30
| 1 | 6 | 233 | 198 | 225 | 375 | 601.0961 | 142.3277 | -- | < 1 | 268 | 197
| Mean |  | 227.6667 | 201.1667 | 225.3333 | 374.1667 | 591.6281 | 142.3113 |  |  | 280 | 200.5
| Median |  | 233 | 201.25 | 225.5 | 374 | 589.8617 | 142.1002 |  |  | 279 | 199
| Std Dev. |  | 14.3898 | 3.1091 | 2.5033 | 14.5797 | 4.798 | 0.83208 |  |  | 8.7636 | 4.0866
| Rel.Std.Dev. |  | 6.321% | 1.546% | 1.111% | 3.897% | 0.811% | 0.5847% |  |  | 3.13% | 2.038%
| PDM3 |  | 6.51% | -5.888% | 5.418% | 75.05% | 176.8% | -33.42% |  |  | 30.99% | -6.2%
| Z-Score (Absolute) |  | 0.07 | 0.48 | 0.07 | 2.46 | 6.15 | 1.48 | NA | NA | 0.86 | 0.49
| Analytical results for Gd in OREAS 465 (Certified Value 584 ppm) | [16 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 599 | 598 | 574 | 540 | 615 | 570.4 | 555.07 | 828.1257 | 633 | 605 | 576.1 | 597 | 604 | 549 |  | 1
| 1 | 2 | 586 | 584 | 590 | 554 | 576 | 574 | 554.26 | 826.2213 | 616 | 606 | 555.3 | 738 | 590 | 532 |  | 8
| 1 | 3 | 599 | 609 | 564 | 554 | 616 | 569.5 | 539.07 | 828.2898 | 607 | 606 | 581.3 | 723 | 588 | 539 |  | 16
| 1 | 4 | 595 | 581 | 658 | 560 | 587 | 572.6 | 562.7 | 824.8291 | 607 | 602.5 | 571.8 | 580 | 670 | 544 |  | 584.393
| 1 | 5 | 593 | 598 | 636 | 547 | 586 | 565.2 | 549.77 | 828.4186 | 616 | 638 | 567.4 | 729 | 592 | 532 |  | 31
| 1 | 6 | 598 | 587 | 660 | 542 | 607 | 570.5 | 576.13 | 825.1736 | 607 | 601 | 550.9 | 717 | 648 | 544
| Mean |  | 595 | 592.8333 | 613.6667 | 549.5 | 597.8333 | 570.3667 | 556.1667 | 826.843 | 614.3333 | 609.75 | 567.1333 | 680.6667 | 615.3333 | 540
| Median |  | 596.5 | 592.5 | 613 | 550.5 | 597 | 570.45 | 554.665 | 827.1735 | 611.5 | 605.5 | 569.6 | 720 | 598 | 541.5
| Std Dev. |  | 5.02 | 10.6474 | 42.9216 | 7.791 | 16.9873 | 3.0217 | 12.4842 | 1.6401 | 10.1522 | 13.9848 | 11.8869 | 71.9296 | 34.9781 | 6.957
| Rel.Std.Dev. |  | 0.8437% | 1.796% | 6.994% | 1.418% | 2.841% | 0.5298% | 2.245% | 0.1984% | 1.653% | 2.294% | 2.096% | 10.57% | 5.684% | 1.288%
| PDM3 |  | 1.815% | 1.444% | 5.009% | -5.971% | 2.3% | -2.4% | -4.83% | 41.49% | 5.123% | 4.339% | -2.953% | 16.47% | 5.294% | -7.596%
| Z-Score (Absolute) |  | 0.04 | 0.11 | 0.52 | 1.41 | 0.04 | 0.78 | 1.21 | 6.91 | 0.43 | 0.23 | 0.88 | 2.53 | 0.57 | 1.69
| Analytical results for Ge in OREAS 465 (Indicative Value 60 ppm) | [16 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 20 | 22 | 23 | [12 empty] | ppm
|  |  | BF*MS | PF*MS | PF*OES/MS | [12 empty] | 0
0
| 1 | 1 | < 1 | 60 | < 1 | [12 empty] | 1
| 1 | 2 | < 1 | 60 | < 1 | [12 empty] | 21
| 1 | 3 | < 1 | 60 | < 1 | [12 empty] | 16
| 1 | 4 | < 1 | 60 | < 1 | [12 empty] | 60
| 1 | 5 | < 1 | 60 | < 1 | [12 empty] | 27
| 1 | 6 | < 1 | 60 | < 1
| Mean |  |  | 60
| Median |  |  | 60
| Std Dev. |  |  | 0
| Rel.Std.Dev. |  |  | 0%
| PDM3 |  |  | 0%
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for Hf in OREAS 465 (Certified Value 41.4 ppm) | [16 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 20 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*OES/MS |  |  |  |  | 1
1
| 1 | 1 | 42.6 | 44.5 | 49 | 45 | 47 | 21.6 | 26.6 | N.A. | 13 | 27 | 33 |  |  |  |  | 1
| 1 | 2 | 43.2 | 44.8 | 49 | 45.6 | 44.2 | 25.5 | 28.8 | N.A. | 16 | 46 | 37 |  |  |  |  | 9
| 1 | 3 | 44.5 | 44.9 | 48 | 46.3 | 45.7 | 23 | 28.3 | N.A. | 13 | 46 | 37 |  |  |  |  | 16
| 1 | 4 | 44.4 | 47.6 | 46 | 46 | 43.8 | 23.6 | 25.6 | N.A. | 14 | 27 | 38 |  |  |  |  | 41.4062
| 1 | 5 | 45 | 43 | 47 | 46.2 | 43.4 | 15.7 | 25.7 | N.A. | 18 | 47 | 37 |  |  |  |  | 32
| 1 | 6 | 45 | 43.8 | 48 | 44.6 | 44.8 | 27.9 | 28.6 | -- | 15 | 46 | 36
| Mean |  | 44.1167 | 44.7667 | 47.8333 | 45.6167 | 44.8167 | 22.8833 | 27.2667 |  | 14.8333 | 39.8333 | 36.3333
| Median |  | 44.45 | 44.65 | 48 | 45.8 | 44.5 | 23.3 | 27.45 |  | 14.5 | 46 | 37
| Std Dev. |  | 0.99281 | 1.5603 | 1.169 | 0.68823 | 1.3393 | 4.1422 | 1.4747 |  | 1.9408 | 9.9482 | 1.7512
| Rel.Std.Dev. |  | 2.25% | 3.485% | 2.444% | 1.509% | 2.988% | 18.1% | 5.408% |  | 13.08% | 24.97% | 4.82%
| PDM3 |  | 6.546% | 8.116% | 15.52% | 10.17% | 8.236% | -44.73% | -34.15% |  | -64.18% | -3.799% | -12.25%
| Z-Score (Absolute) |  | 0.34 | 0.44 | 0.92 | 0.57 | 0.44 | 2.99 | 2.3 | NA | 4.25 | 0.34 | 0.78
| Analytical results for Ho in OREAS 465 (Certified Value 27.7 ppm) | [16 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 26.5 | 27.5 | 29.3 | 28.1 | 28.4 | 29.116 | 28.1 | 27.3 | 18.7978 | < 44 | 28.1 | 22.5419 | 24.4 | 29.8 | 24.2 | 1
| 1 | 2 | 27.6 | 26.8 | 30.4 | 28.4 | 26.8 | 29.4001 | 28.6 | 26.4 | 19.7814 | < 44 | 27.4 | 23.7691 | 30.2 | 28.2 | 24.6 | 10
| 1 | 3 | 27.4 | 28.8 | 28.9 | 28.8 | 28.4 | 29.8365 | 28 | 26.6 | 19.2251 | < 44 | 27.9 | 23.9365 | 30.4 | 29.4 | 24.9 | 16
| 1 | 4 | 26.1 | 27.5 | 31.1 | 28.8 | 27.2 | 29.2084 | 28 | 26.8 | 19.5959 | < 44 | 27.3 | 23.8152 | 24 | 33.6 | 24.8 | 27.7013
| 1 | 5 | 27.2 | 27.4 | 31.1 | 28.3 | 27.1 | 28.7582 | 27.9 | 27.1 | 19.3254 | < 44 | 29.2 | 23.8894 | 30.7 | 29.6 | 24.2 | 33
| 1 | 6 | 25.7 | 27.8 | 31.8 | 28.5 | 28.3 | 29.7279 | 28.1 | 27 | 19.0748 | < 44 | 27.2 | 24.3021 | 29.9 | 32 | 24.5
| Mean |  | 26.75 | 27.6333 | 30.4333 | 28.4833 | 27.7 | 29.3412 | 28.1167 | 26.8667 | 19.3001 |  | 27.85 | 23.709 | 28.2667 | 30.4333 | 24.5333
| Median |  | 26.85 | 27.5 | 30.75 | 28.45 | 27.75 | 29.3043 | 28.05 | 26.9 | 19.2752 |  | 27.65 | 23.8523 | 30.05 | 29.7 | 24.55
| Std Dev. |  | 0.76616 | 0.65929 | 1.1308 | 0.27869 | 0.74297 | 0.40166 | 0.24833 | 0.33267 | 0.35454 |  | 0.75033 | 0.60217 | 3.1633 | 1.9816 | 0.29439
| Rel.Std.Dev. |  | 2.864% | 2.386% | 3.716% | 0.9784% | 2.682% | 1.369% | 0.8832% | 1.238% | 1.837% |  | 2.694% | 2.54% | 11.19% | 6.511% | 1.2%
| PDM3 |  | -3.434% | -0.2454% | 9.862% | 2.823% | -0.004712% | 5.92% | 1.499% | -3.013% | -30.33% |  | 0.5368% | -14.41% | 2.041% | 9.862% | -11.44%
| Z-Score (Absolute) |  | 0.71 | 0.26 | 1.85 | 0.49 | 0.05 | 1.09 | 0.24 | 0.63 | 5.91 | NA | 0.05 | 2.67 | 0.34 | 1.85 | 2.26
| Analytical results for In in OREAS 465 (Certified Value 3.47 ppm) | [16 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
3
| 1 | 1 | 3.5391 | 3.7 | 3 | 2.3 | 3.4 | < 0.2 | [9 empty] | 1
| 1 | 2 | 3.4847 | 3.5 | 3 | 2.8 | 3.6 | 3.6 | [9 empty] | 0
| 1 | 3 | 3.7187 | 3.8 | 3.5 | 2.8 | 3.2 | 3.6 | [9 empty] | 16
| 1 | 4 | 3.5094 | 3.4 | 3 | 2.3 | 3.8 | 3.7 | [9 empty] | 3.4687
| 1 | 5 | 3.352 | 3.5 | 3.5 | 2.9 | 3.4 | 3.6 | [9 empty] | 34
| 1 | 6 | 3.6373 | 3.3 | 3.5 | 2.8 | 3.6 | 3.1
| Mean |  | 3.5402 | 3.5333 | 3.25 | 2.65 | 3.5 | 3.52
| Median |  | 3.5243 | 3.5 | 3.25 | 2.8 | 3.5 | 3.6
| Std Dev. |  | 0.12707 | 0.18619 | 0.27386 | 0.27386 | 0.20976 | 0.23875
| Rel.Std.Dev. |  | 3.589% | 5.27% | 8.427% | 10.33% | 5.993% | 6.783%
| PDM3 |  | 2.061% | 1.863% | -6.305% | -23.6% | 0.9021% | 1.479%
| Z-Score (Absolute) |  | 0.76 | 0.59 | 6.55 | 21.66 | 0.25 | 0.25
| Analytical results for K in OREAS 465 (Indicative Value 0.099 wt.%) | [16 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.229 | 0.06 | < 0.01 | 0.19246 | N.A. | < 0.1 | 0.01254 | 0.03 | < 0.02 | < 0.1 | [5 empty] | 1
| 1 | 2 | 0.234 | 0.06 | < 0.01 | 0.19636 | N.A. | < 0.1 | 0.01478 | 0.02 | < 0.02 | < 0.1 | [5 empty] | 11
| 1 | 3 | 0.231 | 0.06 | < 0.01 | 0.1975 | N.A. | < 0.1 | 0.01212 | 0.02 | < 0.02 | < 0.1 | [5 empty] | 16
| 1 | 4 | 0.223 | < 0.05 | < 0.01 | 0.18782 | N.A. | < 0.1 | 0.01129 | 0.02 | < 0.02 | < 0.1 | [5 empty] | 0.099038
| 1 | 5 | 0.214 | < 0.05 | < 0.01 | 0.20145 | N.A. | < 0.1 | 0.01112 | 0.02 | < 0.02 | < 0.1 | [5 empty] | 21
| 1 | 6 | 0.21 | < 0.05 | < 0.01 | 0.20842 | -- | < 0.1 | 0.0093 | 0.02 | < 0.02 | < 0.1
| Mean |  | 0.2235 | 0.06 |  | 0.19733 |  |  | 0.011858 | 0.021667
| Median |  | 0.226 | 0.06 |  | 0.19693 |  |  | 0.011705 | 0.02
| Std Dev. |  | 0.0096902 | 0 |  | 0.0071433 |  |  | 0.0018156 | 0.0040825
| Rel.Std.Dev. |  | 4.336% | 0% |  | 3.62% |  |  | 15.31% | 18.84%
| PDM3 |  | 125.7% | -39.42% |  | 99.25% |  |  | -88.03% | -78.12%
| Z-Score (Absolute) |  | 3.98 | 0 | NA | 3.41 | NA | NA | 0.67 | 0.5 | NA | NA
| Analytical results for La in OREAS 465 (Certified Value 2.41 wt.%) | [16 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*OES | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 2.49 | 2.37 | 2.34 | > 1 | 2.43 | 2.1261 | 2.4458 | 2.4069 | 2.3364 | 2.285 | 2.3538 | 2.4119 | 2.39 | 2.5 | 2.51 | 1
| 1 | 2 | 2.55 | 2.23 | 2.34 | > 1 | 2.42 | 2.1566 | 2.4543 | 2.4065 | 2.396 | 2.234 | 2.3417 | 2.518 | 2.38 | 2.37 | 2.45 | 0
| 1 | 3 | 2.52 | 2.43 | 2.34 | > 1 | 2.48 | 2.1686 | 2.3862 | 2.4053 | 2.3646 | 2.26 | 2.3485 | 2.5406 | 2.39 | 2.4 | 2.38 | 16
| 1 | 4 | 2.41 | 2.38 | 2.38 | > 1 | 2.39 | 2.1623 | 2.3908 | 2.4236 | 2.3536 | 2.268 | 2.3452 | 2.5394 | 2.46 | 2.64 | 2.55 | 2.4128
| 1 | 5 | 2.52 | 2.36 | 2.37 | > 1 | 2.44 | 2.1619 | 2.3671 | 2.417 | 2.3954 | 2.26 | 2.4599 | 2.5303 | 2.44 | 2.43 | 2.54 | 35
| 1 | 6 | 2.43 | 2.41 | 2.33 | > 1 | 2.52 | 2.2127 | 2.4143 | 2.4424 | 2.3728 | 2.242 | 2.3209 | 2.5901 | 2.39 | 2.51 | 2.39
| Mean |  | 2.4867 | 2.3633 | 2.35 |  | 2.4467 | 2.1647 | 2.4098 | 2.4169 | 2.3698 | 2.2582 | 2.3617 | 2.5217 | 2.4083 | 2.475 | 2.47
| Median |  | 2.505 | 2.375 | 2.34 |  | 2.435 | 2.1621 | 2.4026 | 2.412 | 2.3687 | 2.26 | 2.3468 | 2.5348 | 2.39 | 2.465 | 2.48
| Std Dev. |  | 0.055377 | 0.070333 | 0.02 |  | 0.046332 | 0.027882 | 0.034735 | 0.014407 | 0.023472 | 0.018269 | 0.04943 | 0.059153 | 0.033116 | 0.097724 | 0.074565
| Rel.Std.Dev. |  | 2.227% | 2.976% | 0.8511% |  | 1.894% | 1.288% | 1.441% | 0.5961% | 0.9905% | 0.809% | 2.093% | 2.346% | 1.375% | 3.948% | 3.019%
| PDM3 |  | 3.059% | -2.052% | -2.605% |  | 1.402% | -10.28% | -0.1276% | 0.17% | -1.784% | -6.411% | -2.121% | 4.512% | -0.1871% | 2.576% | 2.369%
| Z-Score (Absolute) |  | 0.87 | 0.21 | 0.66 | NA | 0.42 | 2.75 | 0.01 | 0.09 | 0.44 | 1.7 | 0.75 | 1.51 | 0.01 | 0.74 | 0.68
| Analytical results for Li in OREAS 465 (Indicative Value 3.33 ppm) | [16 empty] | Indicative
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | PF*OES | PF*OES | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 5 | < 5 | 3 | < 10 | 4 | < 10 | [9 empty] | 1
| 1 | 2 | < 5 | < 5 | 3 | < 10 | 3 | < 10 | [9 empty] | 16
| 1 | 3 | < 5 | < 5 | 3 | < 10 | 4 | < 10 | [9 empty] | 16
| 1 | 4 | < 5 | < 5 | 3 | < 10 | 4 | < 10 | [9 empty] | 3.3333
| 1 | 5 | < 5 | < 5 | 3 | < 10 | 3 | < 10 | [9 empty] | 22
| 1 | 6 | < 5 | < 5 | 4 | < 10 | 3 | < 10
| Mean |  |  |  | 3.1667 |  | 3.5
| Median |  |  |  | 3 |  | 3.5
| Std Dev. |  |  |  | 0.40825 |  | 0.54772
| Rel.Std.Dev. |  |  |  | 12.89% |  | 15.65%
| PDM3 |  |  |  | -5% |  | 5%
| Z-Score (Absolute) |  | NA | NA | 0.67 | NA | 0.67 | NA
| Analytical results for Lu in OREAS 465 (Certified Value 2.1 ppm) | [16 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  | 2
3
| 1 | 1 | 2.05 | 2.17 | 2.08 | 2.1 | 2.09 | 2.1661 | 2 | 2 | < 44 | 2 | 1.6265 | 1.03 | 2.2 | 2.66 |  | 1
| 1 | 2 | 2.02 | 2.01 | 2.12 | 2.05 | 2.08 | 2.3039 | 2.1 | 1.9 | < 44 | 2 | 1.6794 | 1.19 | 1.8 | 2.06 |  | 1
| 1 | 3 | 2.13 | 2.17 | 2.09 | 2.19 | 2.17 | 2.2135 | 1.9 | 1.9 | < 44 | 2.5 | 1.6862 | 1.15 | 2 | 2.17 |  | 16
| 1 | 4 | 2.15 | 2.01 | 2.15 | 2.15 | 2.05 | 2.0135 | 2.1 | 1.9 | < 44 | 2 | 1.6736 | 0.96 | 2.4 | 2.13 |  | 2.1042
| 1 | 5 | 2.17 | 2.12 | 2.11 | 2.1 | 2.1 | 2.2429 | 2 | 2 | < 44 | 2.5 | 1.6614 | 1.21 | 2.2 | 2.06 |  | 36
| 1 | 6 | 2.14 | 2.04 | 2.27 | 2.13 | 2.1 | 2.1545 | 1.9 | 2.1 | < 44 | 2 | 1.6875 | 1.17 | 2.2 | 8.21
| Mean |  | 2.11 | 2.0867 | 2.1367 | 2.12 | 2.0983 | 2.1824 | 2 | 1.9667 |  | 2.1667 | 1.6691 | 1.1183 | 2.1333 | 3.215
| Median |  | 2.135 | 2.08 | 2.115 | 2.115 | 2.095 | 2.1898 | 2 | 1.95 |  | 2 | 1.6765 | 1.16 | 2.2 | 2.15
| Std Dev. |  | 0.060332 | 0.07607 | 0.069761 | 0.048166 | 0.039707 | 0.098944 | 0.089443 | 0.08165 |  | 0.2582 | 0.022951 | 0.10008 | 0.20656 | 2.4575
| Rel.Std.Dev. |  | 2.859% | 3.646% | 3.265% | 2.272% | 1.892% | 4.534% | 4.472% | 4.152% |  | 11.92% | 1.375% | 8.949% | 9.682% | 76.44%
| PDM3 |  | 0.2776% | -0.8313% | 1.545% | 0.7529% | -0.2768% | 3.719% | -4.95% | -6.534% |  | 2.971% | -20.68% | -46.85% | 1.387% | 52.79%
| Z-Score (Absolute) |  | 0.05 | 0.2 | 0.05 | 0.16 | 0.07 | 0.85 | 1.15 | 1.51 | NA | 0.67 | 4.77 | 10.79 | 1.04 | 0
| Analytical results for Mg in OREAS 465 (Certified Value 0.392 wt.%) | [16 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 3
3
| 1 | 1 | 0.39794 | 0.41 | 0.41 | 0.38519 | N.A. | 0.425 | 0.36642 | 0.32 | 0.39 | 0.42 | 0.36 |  |  |  |  | 1
| 1 | 2 | 0.40104 | 0.41 | 0.41 | 0.37986 | N.A. | 0.415 | 0.36377 | 0.38 | 0.39 | 0.4 | 0.36 |  |  |  |  | 2
| 1 | 3 | 0.39691 | 0.41 | 0.41 | 0.39675 | N.A. | 0.42 | 0.37306 | 0.38 | 0.4 | 0.41 | 0.35 |  |  |  |  | 16
| 1 | 4 | 0.39381 | 0.41 | 0.41 | 0.38126 | N.A. | 0.41 | 0.36817 | 0.31 | 0.39 | 0.41 | 0.35 |  |  |  |  | 0.39174
| 1 | 5 | 0.38967 | 0.42 | 0.41 | 0.3866 | N.A. | 0.42 | 0.37233 | 0.38 | 0.38 | 0.39 | 0.37 |  |  |  |  | 37
| 1 | 6 | 0.39174 | 0.41 | 0.41 | 0.38511 | -- | 0.42 | 0.36491 | 0.38 | 0.38 | 0.4 | 0.34
| Mean |  | 0.39518 | 0.41167 | 0.41 | 0.3858 |  | 0.41833 | 0.36811 | 0.35833 | 0.38833 | 0.405 | 0.355
| Median |  | 0.39536 | 0.41 | 0.41 | 0.38515 |  | 0.42 | 0.36729 | 0.38 | 0.39 | 0.405 | 0.355
| Std Dev. |  | 0.0042197 | 0.0040825 | 0 | 0.0059531 |  | 0.005164 | 0.0038535 | 0.033714 | 0.0075277 | 0.010488 | 0.010488
| Rel.Std.Dev. |  | 1.068% | 0.9917% | 0% | 1.543% |  | 1.234% | 1.047% | 9.409% | 1.938% | 2.59% | 2.954%
| PDM3 |  | 0.8788% | 5.086% | 4.661% | -1.518% |  | 6.788% | -6.033% | -8.528% | -0.8702% | 3.384% | -9.379%
| Z-Score (Absolute) |  | 0.15 | 0.85 | 0.78 | 0.26 | NA | 1.14 | 1.01 | 0.5 | 0.15 | 0.57 | 1.57
| Analytical results for Mn in OREAS 465 (Certified Value 0.263 wt.%) | [16 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | [5 empty] | 3
3
| 1 | 1 | 0.3 | 0.26 | 0.26849 | N.A. | 0.29 | 0.23923 | 0.07 | 0.273 | 0.26 | 0.234 | [5 empty] | 1
| 1 | 2 | 0.3 | 0.26 | 0.26135 | N.A. | 0.285 | 0.2359 | 0.25 | 0.269 | 0.254 | 0.234 | [5 empty] | 3
| 1 | 3 | 0.3 | 0.26 | 0.26807 | N.A. | 0.285 | 0.24357 | 0.25 | 0.27 | 0.258 | 0.24 | [5 empty] | 16
| 1 | 4 | 0.3 | 0.27 | 0.26544 | N.A. | 0.29 | 0.24063 | 0.06 | 0.272 | 0.258 | 0.233 | [5 empty] | 0.26252
| 1 | 5 | 0.3 | 0.26 | 0.26514 | N.A. | 0.29 | 0.24164 | 0.25 | 0.271 | 0.254 | 0.232 | [5 empty] | 38
| 1 | 6 | 0.2 | 0.28 | 0.27198 | -- | 0.29 | 0.23567 | 0.24 | 0.267 | 0.252 | 0.231
| Mean |  | 0.28333 | 0.265 | 0.26675 |  | 0.28833 | 0.23944 | 0.18667 | 0.27033 | 0.256 | 0.234
| Median |  | 0.3 | 0.26 | 0.26676 |  | 0.29 | 0.23993 | 0.245 | 0.2705 | 0.256 | 0.2335
| Std Dev. |  | 0.040825 | 0.0083666 | 0.0036209 |  | 0.002582 | 0.0031655 | 0.094375 | 0.0021602 | 0.0030984 | 0.0031623
| Rel.Std.Dev. |  | 14.41% | 3.157% | 1.357% |  | 0.8955% | 1.322% | 50.56% | 0.7991% | 1.21% | 1.351%
| PDM3 |  | 7.927% | 0.9434% | 1.608% |  | 9.831% | -8.793% | -28.9% | 2.975% | -2.485% | -10.87%
| Z-Score (Absolute) |  | 0.67 | 0 | 0.15 | NA | 0.83 | 0.71 | 2.38 | 0.26 | 0.19 | 0.89
| Analytical results for Mo in OREAS 465 (Certified Value 114 ppm) | [16 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [7 empty] | ppm
|  |  | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 112 | 61.4969 | N.A. | 120 | 105.3972 | 103 | 120 | 118 | [7 empty] | 1
| 1 | 2 | 112 | 60.7224 | N.A. | 120 | 112.2667 | 109 | 110 | 116 | [7 empty] | 0
| 1 | 3 | 112 | 59.9124 | N.A. | 115 | 112.4797 | 108 | 115 | 115 | [7 empty] | 16
| 1 | 4 | 110 | 60.0057 | N.A. | 120 | 112.4673 | 101 | 125 | 117 | [7 empty] | 113.945
| 1 | 5 | 123 | 59.527 | N.A. | 125 | 114.0735 | 108 | 115 | 112 | [7 empty] | 39
| 1 | 6 | 115 | 60.3904 | -- | 120 | 124.869 | 108 | 115 | 115
| Mean |  | 114 | 60.3425 |  | 120 | 113.5923 | 106.1667 | 116.6667 | 115.5
| Median |  | 112 | 60.1981 |  | 120 | 112.4735 | 108 | 115 | 115.5
| Std Dev. |  | 4.6904 | 0.69899 |  | 3.1623 | 6.3058 | 3.3116 | 5.164 | 2.0736
| Rel.Std.Dev. |  | 4.114% | 1.158% |  | 2.635% | 5.551% | 3.119% | 4.426% | 1.795%
| PDM3 |  | 0.04824% | -47.04% |  | 5.314% | -0.3096% | -6.826% | 2.389% | 1.365%
| Z-Score (Absolute) |  | 0 | 13.57 | NA | 1.52 | 0.67 | 1.98 | 0.67 | 0.38
| Analytical results for Na in OREAS 465 (Indicative Value 0.078 wt.%) | [16 empty] | Indicative
Na | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 11 | 14 | 18 | 20 | [11 empty] | wt.%
|  |  | BF*OES | BF*OES | BF*OES | BF*OES | [11 empty] | 3
3
| 1 | 1 | 0.08 | N.A. | 0.076706 | 0.61 | [11 empty] | 1
| 1 | 2 | 0.07 | N.A. | 0.083605 | 0.57 | [11 empty] | 4
| 1 | 3 | 0.07 | N.A. | 0.086869 | 0.52 | [11 empty] | 16
| 1 | 4 | 0.08 | N.A. | 0.084421 | 0.56 | [11 empty] | 0.077808
| 1 | 5 | 0.07 | N.A. | 0.08405 | 0.51 | [11 empty] | 23
| 1 | 6 | 0.07 | -- | 0.078042 | 0.54
| Mean |  | 0.073333 |  | 0.082282 | 0.55167
| Median |  | 0.07 |  | 0.083828 | 0.55
| Std Dev. |  | 0.005164 |  | 0.0039893 | 0.03656
| Rel.Std.Dev. |  | 7.042% |  | 4.848% | 6.627%
| PDM3 |  | -5.751% |  | 5.751% | 609%
| Z-Score (Absolute) |  | 0.67 | NA | 0 | 35.37
| Analytical results for Nb in OREAS 465 (Certified Value 4680 ppm) | [16 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*OES | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 4990 | 4700 | 4670 | > 2500 | 4540 | 4644 | 4589 | N.A. | 2050 | 4211.7553 | 1250 | 5050 | 5180 |  |  | 1
| 1 | 2 | > 2500 | 4310 | 4570 | > 2500 | 4530 | 4644 | 4443 | N.A. | 2315 | 3987.8005 | 3690 | 5140 | 5150 |  |  | 27
| 1 | 3 | 4930 | 4710 | 4690 | > 2500 | 4580 | 4623 | 4466 | N.A. | 1850 | 3654.4153 | 3730 | 5090 | 5310 |  |  | 16
| 1 | 4 | 4810 | 4610 | 4160 | > 2500 | 4430 | 4684 | 4488 | N.A. | 2195 | 4356.5609 | 1110 | 4870 | 5020 |  |  | 4680.1729
| 1 | 5 | > 2500 | 4620 | 4130 | > 2500 | 4550 | 4666 | 4490 | N.A. | 2530 | 4513.5654 | 3740 | 4990 | 5150 |  |  | 40
| 1 | 6 | 4860 | 4770 | 4100 | > 2500 | 4700 | 4665 | 4631 | -- | 2125 | 4388.0965 | 3530 | 4740 | 5130
| Mean |  | 4897.5 | 4620 | 4386.6667 |  | 4555 | 4654.3333 | 4517.8333 |  | 2177.5 | 4185.3656 | 2841.6667 | 4980 | 5156.6667
| Median |  | 4895 | 4660 | 4365 |  | 4545 | 4654.5 | 4489 |  | 2160 | 4284.1581 | 3610 | 5020 | 5150
| Std Dev. |  | 78.8987 | 163.2176 | 284.7221 |  | 87.3499 | 21.5654 | 74.596 |  | 232.266 | 316.1517 | 1290.0917 | 149.9333 | 93.3095
| Rel.Std.Dev. |  | 1.611% | 3.533% | 6.491% |  | 1.918% | 0.4633% | 1.651% |  | 10.67% | 7.554% | 45.4% | 3.011% | 1.809%
| PDM3 |  | 4.644% | -1.286% | -6.271% |  | -2.675% | -0.5521% | -3.469% |  | -53.47% | -10.57% | -39.28% | 6.406% | 10.18%
| Z-Score (Absolute) |  | 0.88 | 0.33 | 0.43 | NA | 0 | 0.25 | 0.1 | NA | 6.09 | 0.67 | 4.39 | 1.09 | 1.54
| Analytical results for Nd in OREAS 465 (Certified Value 1.18 wt.%) | [16 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | wt.%
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 1.22 | 1.2 | 1.19 | > 1 | 1.19 | 1.1649 | 1.1647 | 1.1205 | 0.9679 | 1.175 | 1.1751 | 1.0584 | 0.942 | 1.25 | 1.24 | 1
| 1 | 2 | 1.245 | 1.175 | 1.2 | > 1 | 1.205 | 1.199 | 1.1789 | 1.1175 | 0.97493 | 1.149 | 1.1655 | 1.0681 | 1.16 | 1.17 | 1.23 | 0
| 1 | 3 | 1.25 | 1.185 | 1.195 | > 1 | 1.19 | 1.2048 | 1.1534 | 1.1204 | 0.96177 | 1.149 | 1.164 | 1.0919 | 1.16 | 1.21 | 1.25 | 16
| 1 | 4 | 1.185 | 1.205 | 1.145 | > 1 | 1.18 | 1.2053 | 1.153 | 1.1222 | 0.95752 | 1.166 | 1.1584 | 1.0747 | 1.12 | 1.32 | 1.24 | 1.1764
| 1 | 5 | 1.26 | 1.2 | 1.145 | > 1 | 1.23 | 1.2103 | 1.1376 | 1.1119 | 0.95985 | 1.166 | 1.2254 | 1.0805 | 1.19 | 1.22 | 1.23 | 41
| 1 | 6 | 1.195 | 1.205 | 1.13 | > 1 | 1.22 | 1.2249 | 1.1497 | 1.1188 | 0.97656 | 1.149 | 1.1462 | 1.0731 | 1.15 | 1.25 | 1.17
| Mean |  | 1.2258 | 1.195 | 1.1675 |  | 1.2025 | 1.2015 | 1.1562 | 1.1186 | 0.96642 | 1.159 | 1.1724 | 1.0744 | 1.1203 | 1.2367 | 1.2267
| Median |  | 1.2325 | 1.2 | 1.1675 |  | 1.1975 | 1.2051 | 1.1532 | 1.1196 | 0.96484 | 1.1575 | 1.1647 | 1.0739 | 1.155 | 1.235 | 1.235
| Std Dev. |  | 0.03089 | 0.012247 | 0.030781 |  | 0.019429 | 0.019975 | 0.014104 | 0.0036594 | 0.0080145 | 0.011437 | 0.027629 | 0.011323 | 0.090203 | 0.050465 | 0.028752
| Rel.Std.Dev. |  | 2.52% | 1.025% | 2.637% |  | 1.616% | 1.662% | 1.22% | 0.3272% | 0.8293% | 0.9868% | 2.357% | 1.054% | 8.051% | 4.081% | 2.344%
| PDM3 |  | 4.204% | 1.583% | -0.7549% |  | 2.22% | 2.138% | -1.713% | -4.915% | -17.85% | -1.477% | -0.3355% | -8.665% | -4.764% | 5.125% | 4.275%
| Z-Score (Absolute) |  | 1.1 | 0.55 | 0.05 | NA | 0.68 | 0.67 | 0.15 | 0.83 | 3.58 | 0.11 | 0.05 | 1.63 | 0.16 | 1.3 | 1.32
| Analytical results for Ni in OREAS 465 (Indicative Value 101 ppm) | [16 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | BF*OES | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 102 | N.A. | 100 | 114.6768 | 27 | 100 | 150 | 99 | [7 empty] | 1
| 1 | 2 | 96 | N.A. | 100 | 128.6745 | 102 | 90 | 100 | 97 | [7 empty] | 6
| 1 | 3 | 99 | N.A. | 100 | 122.4979 | 103 | 110 | 100 | 101 | [7 empty] | 16
| 1 | 4 | 65 | N.A. | 100 | 123.9885 | 22 | 110 | 150 | 98 | [7 empty] | 100.7898
| 1 | 5 | 62 | N.A. | 100 | 147.1376 | 94 | 90 | 100 | 93 | [7 empty] | 24
| 1 | 6 | 63 | -- | 100 | 253.207 | 93 | 100 | 100 | 91
| Mean |  | 81.1667 |  | 100 | 148.3637 | 73.5 | 100 | 116.6667 | 96.5
| Median |  | 80.5 |  | 100 | 126.3315 | 93.5 | 100 | 100 | 97.5
| Std Dev. |  | 19.6511 |  | 1.5567e-14 | 52.4962 | 38.2034 | 8.9443 | 25.8199 | 3.7815
| Rel.Std.Dev. |  | 24.21% |  | 1.557e-14% | 35.38% | 51.98% | 8.944% | 22.13% | 3.919%
| PDM3 |  | -19.47% |  | -0.7836% | 47.2% | -27.08% | -0.7836% | 15.75% | -4.256%
| Z-Score (Absolute) |  | 0.78 | NA | 0 | 1.14 | 0.67 | 0 | 0.69 | 0.15
| Analytical results for P in OREAS 465 (Certified Value 3.81 wt.%) | [16 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | [5 empty] | wt.%
|  |  | PF*OES | PF*OES | BF*OES | PF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS | [5 empty] | 2
3
| 1 | 1 | 3.754 | 3.71 | 3.73 | 3.7416 | N.A. | 3.91 | 3.9431 | 3.38 | 4.06 | 2.81 | [5 empty] | 1
| 1 | 2 | 3.799 | 3.75 | 3.68 | 3.7581 | N.A. | 3.78 | 3.817 | 3.36 | 3.86 | 2.68 | [5 empty] | 7
| 1 | 3 | 3.776 | 3.63 | 3.67 | 3.8636 | N.A. | 3.865 | 3.9344 | 3.41 | 4.08 | 2.77 | [5 empty] | 16
| 1 | 4 | 3.724 | 3.72 | 3.61 | 3.8716 | N.A. | 3.835 | 3.9056 | 3.24 | 3.9 | 2.5 | [5 empty] | 3.8122
| 1 | 5 | 3.745 | 3.75 | 3.61 | 3.8008 | N.A. | 3.795 | 3.9505 | 3.33 | 4.12 | 2.52 | [5 empty] | 42
| 1 | 6 | 3.703 | 3.72 | 3.72 | 3.8164 | -- | 3.795 | 3.8685 | 3.44 | 4.04 | 2.62
| Mean |  | 3.7502 | 3.7133 | 3.67 | 3.8087 |  | 3.83 | 3.9032 | 3.36 | 4.01 | 2.65
| Median |  | 3.7495 | 3.72 | 3.675 | 3.8086 |  | 3.815 | 3.92 | 3.37 | 4.05 | 2.65
| Std Dev. |  | 0.034649 | 0.044121 | 0.051769 | 0.053211 |  | 0.0502 | 0.051859 | 0.070143 | 0.10488 | 0.12744
| Rel.Std.Dev. |  | 0.9239% | 1.188% | 1.411% | 1.397% |  | 1.311% | 1.329% | 2.088% | 2.615% | 4.809%
| PDM3 |  | -1.627% | -2.593% | -3.73% | -0.09222% |  | 0.4671% | 2.386% | -11.86% | 5.189% | -30.49%
| Z-Score (Absolute) |  | 0 | 0.31 | 0.67 | 0.49 | NA | 0.67 | 1.29 | 3.28 | 2.19 | 9.25
| Analytical results for Pb in OREAS 465 (Certified Value 506 ppm) | [16 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 20 | 21 | 22 | 23 | [7 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 529 | 579 | 600 | 490.509 | 262 | 500 | 480 | 510 | [7 empty] | 1
| 1 | 2 | 528 | 509 | 700 | 511.7467 | 274 | 500 | 470 | 496 | [7 empty] | 8
| 1 | 3 | 521 | 503 | 700 | 510.2846 | 266 | 500 | 490 | 491 | [7 empty] | 16
| 1 | 4 | 517 | 499 | 700 | 520.3917 | 274 | 500 | 590 | 522 | [7 empty] | 506.1513
| 1 | 5 | 520 | 521 | 700 | 509.7365 | 274 | 500 | 470 | 505 | [7 empty] | 43
| 1 | 6 | 527 | 497 | 700 | 507.979 | 262 | 500 | 570 | 494
| Mean |  | 523.6667 | 518 | 683.3333 | 508.4412 | 268.6667 | 500 | 511.6667 | 503
| Median |  | 524 | 506 | 700 | 510.0106 | 270 | 500 | 485 | 500.5
| Std Dev. |  | 4.9666 | 31.1063 | 40.8248 | 9.804 | 6.0222 | 0 | 53.8207 | 11.7303
| Rel.Std.Dev. |  | 0.9484% | 6.005% | 5.974% | 1.928% | 2.242% | 0% | 10.52% | 2.332%
| PDM3 |  | 3.46% | 2.341% | 35.01% | 0.4524% | -46.92% | -1.215% | 1.09% | -0.6226%
| Z-Score (Absolute) |  | 2.03 | 0.15 | 18.85 | 0.43 | 24.84 | 0.46 | 0.89 | 0.15
| Analytical results for Pr in OREAS 465 (Certified Value 3772 ppm) | [16 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 3940 | 3640 | 3720 | > 1000 | 3770 | 3745.0915 | 3803.3 | 3952 | 3063.3024 | 3675 | 3752 | 3886.6729 | 3060 | 3850 | 3490 | 1
| 1 | 2 | 4030 | 3570 | 3730 | > 1000 | 3820 | 3833.7277 | 3742 | 3940 | 3082.353 | 3625 | 3706 | 4062.8763 | 3750 | 3680 | 3460 | 0
| 1 | 3 | 4020 | 3870 | 3750 | > 1000 | 3780 | 3878.7984 | 3730.2 | 3904.8 | 3049.4558 | 3642 | 3718.5 | 4042.9545 | 3760 | 3760 | 3570 | 16
| 1 | 4 | 3800 | 3710 | 3510 | > 1000 | 3640 | 3835.4871 | 3723.6 | 3979.4 | 2973.5193 | 3658 | 3740 | 4076.2038 | 2980 | 3700 | 3460 | 3771.8124
| 1 | 5 | 4070 | 3740 | 3530 | > 1000 | 3800 | 3815.5471 | 3714 | 3907.5 | 3116.9763 | 3642 | 3919.5 | 4084.8665 | 3830 | 3730 | 3490 | 44
| 1 | 6 | 3860 | 3770 | 3460 | > 1000 | 3840 | 3877.5942 | 3705.9 | 3953.2 | 3032.0345 | 3625 | 3655 | 4169.9195 | 3740 | 3680 | 3460
| Mean |  | 3953.3333 | 3716.6667 | 3616.6667 |  | 3775 | 3831.041 | 3736.5 | 3939.4833 | 3052.9402 | 3644.5 | 3748.5 | 4053.9156 | 3520 | 3733.3333 | 3488.3333
| Median |  | 3980 | 3725 | 3625 |  | 3790 | 3834.6074 | 3726.9 | 3946 | 3056.3791 | 3642 | 3729.25 | 4069.54 | 3745 | 3715 | 3475
| Std Dev. |  | 106.1446 | 104.2433 | 130.1794 |  | 70.9225 | 49.1982 | 35.0463 | 28.8593 | 48.6415 | 19.4191 | 90.3039 | 92.8247 | 389.4098 | 65.0128 | 42.6224
| Rel.Std.Dev. |  | 2.685% | 2.805% | 3.599% |  | 1.879% | 1.284% | 0.9379% | 0.7326% | 1.593% | 0.5328% | 2.409% | 2.29% | 11.06% | 1.741% | 1.222%
| PDM3 |  | 4.813% | -1.462% | -4.113% |  | 0.08451% | 1.57% | -0.9362% | 4.445% | -19.06% | -3.375% | -0.6181% | 7.479% | -6.676% | -1.02% | -7.516%
| Z-Score (Absolute) |  | 1.58 | 0.13 | 0.86 | NA | 0.49 | 0.7 | 0.01 | 1.48 | 4.93 | 0.65 | 0.15 | 2.55 | 0.25 | 0.01 | 1.78
| Analytical results for Rb in OREAS 465 (Indicative Value 0.98 ppm) | [16 empty] | Indicative
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  | 2
2
| 1 | 1 | 1.4 | 0.5 | 0.6 | 0.4 | 1 | 0.67795 | 0.8 | 1.4 | 3.2296 | N.A. | 2 | 0.4 | 2 | 1.8 |  | 1
| 1 | 2 | 0.9 | 0.6 | 0.6 | 0.4 | < 1 | 0.79964 | 0.9 | 1.2 | 3.2222 | N.A. | 2 | 0.3 | 2 | 1.1 |  | 19
| 1 | 3 | 1.9 | 0.6 | 0.8 | 0.5 | < 1 | 0.83384 | 0.6 | 1.4 | 3.1821 | N.A. | 2 | 0.3 | 1 | 1.3 |  | 16
| 1 | 4 | 0.3 | 0.7 | 0.8 | 0.5 | < 1 | 0.85764 | 0.9 | 1.4 | 3.1785 | N.A. | 2 | 0.4 | 1 | 1.8 |  | 0.98192
| 1 | 5 | 0.5 | 0.6 | 0.8 | 0.4 | < 1 | 0.69061 | 0.8 | 2 | 3.2671 | N.A. | 2 | 0.4 | 1 | 1.1 |  | 25
| 1 | 6 | 0.3 | 0.6 | 0.7 | 0.5 | < 1 | 0.77829 | < 0.5 | 1.4 | 3.178 | -- | 2 | 0.5 | 1 | 1.8
| Mean |  | 0.88333 | 0.6 | 0.71667 | 0.45 | 1 | 0.773 | 0.8 | 1.4667 | 3.2096 |  | 2 | 0.38333 | 1.3333 | 1.4833
| Median |  | 0.7 | 0.6 | 0.75 | 0.45 | 1 | 0.78897 | 0.8 | 1.4 | 3.2022 |  | 2 | 0.4 | 1 | 1.55
| Std Dev. |  | 0.65243 | 0.063246 | 0.098319 | 0.054772 |  | 0.074063 | 0.12247 | 0.27325 | 0.036276 |  | 0 | 0.075277 | 0.5164 | 0.35449
| Rel.Std.Dev. |  | 73.86% | 10.54% | 13.72% | 12.17% |  | 9.581% | 15.31% | 18.63% | 1.13% |  | 0% | 19.64% | 38.73% | 23.9%
| PDM3 |  | -10.04% | -38.89% | -27.01% | -54.17% | 1.842% | -21.28% | -18.53% | 49.37% | 226.9% |  | 103.7% | -60.96% | 35.79% | 51.07%
| Z-Score (Absolute) |  | 0 | 0.44 | 0.26 | 0.67 | 0.18 | 0.17 | 0.13 | 0.74 | 3.62 | NA | 1.74 | 0.78 | 0.7 | 0.93
| Analytical results for Re in OREAS 465 (Indicative Value < 0.1 ppm) | [16 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | [12 empty] | 1
| No | No | 09 | 17 | 22 | [12 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | [12 empty] | 2
2
| 1 | 1 | < 0.1 | < 0.1 | < 1 | [12 empty] | 1
| 1 | 2 | 0.5 | < 0.1 | < 1 | [12 empty] | 20
| 1 | 3 | < 0.1 | < 0.1 | < 1 | [12 empty] | 16
| 1 | 4 | < 0.1 | < 0.1 | < 1 | [12 empty] | < 0.1
| 1 | 5 | < 0.1 | < 0.1 | < 1 | [12 empty] | 26
| 1 | 6 | < 0.1 | < 0.1 | < 1
| Mean |  | 0.5
| Median |  | 0.5
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA
| Analytical results for S in OREAS 465 (Certified Value 1941 ppm) | [16 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 08 | 09 | 17 | 21 | 22 | [10 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | [10 empty] | 0
0
| 1 | 1 | 1820 | 1900 | 2100 | 1800 | 2200 | [10 empty] | 1
| 1 | 2 | 1930 | 1900 | 2000 | 2000 | 2200 | [10 empty] | 11
| 1 | 3 | 1850 | 1800 | 2000 | 1800 | 2000 | [10 empty] | 16
| 1 | 4 | 1760 | 1900 | 2100 | 1800 | 2200 | [10 empty] | 1941
| 1 | 5 | 1860 | 1900 | 2100 | 1900 | 2000 | [10 empty] | 45
| 1 | 6 | 1810 | 1800 | 2000 | 1800 | 2000
| Mean |  | 1838.3333 | 1866.6667 | 2050 | 1850 | 2100
| Median |  | 1835 | 1900 | 2050 | 1800 | 2100
| Std Dev. |  | 57.0672 | 51.6398 | 54.7723 | 83.666 | 109.5445
| Rel.Std.Dev. |  | 3.104% | 2.766% | 2.672% | 4.522% | 5.216%
| PDM3 |  | -5.289% | -3.83% | 5.616% | -4.688% | 8.192%
| Z-Score (Absolute) |  | 0.67 | 0 | 4.36 | 0.4 | 5.55
| Analytical results for Sb in OREAS 465 (Indicative Value 0.98 ppm) | [16 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 11 | 17 | 18 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*OES/MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 1.2 | 2.9 | 1 | 0.99213 | < 1 | 1.1 | [9 empty] | 1
| 1 | 2 | 1.2 | 2.7 | 1 | 0.93243 | < 1 | 1.1 | [9 empty] | 21
| 1 | 3 | 1.4 | 2.9 | 1 | 0.95948 | 1 | 1.2 | [9 empty] | 16
| 1 | 4 | 1.4 | 3.2 | 1 | 0.98606 | < 1 | 1 | [9 empty] | 0.98407
| 1 | 5 | 1.3 | 3 | 1 | 1.213 | < 1 | 1.1 | [9 empty] | 27
| 1 | 6 | 1.5 | 3.1 | 1 | 1.0649 | < 1 | 1.1
| Mean |  | 1.3333 | 2.9667 | 1 | 1.0247 | 1 | 1.1
| Median |  | 1.35 | 2.95 | 1 | 0.9891 | 1 | 1.1
| Std Dev. |  | 0.12111 | 0.17512 | 0 | 0.10236 |  | 0.063246
| Rel.Std.Dev. |  | 9.083% | 5.903% | 0% | 9.99% |  | 5.75%
| PDM3 |  | 35.49% | 201.5% | 1.619% | 4.126% | 1.619% | 11.78%
| Z-Score (Absolute) |  | 1.1 | 7.46 | 0.19 | 0.25 | 2.14 | 0.19
| Analytical results for Sc in OREAS 465 (Indicative Value 171 ppm) | [16 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | [11 empty] | 1
| No | No | 09 | 17 | 22 | 23 | [11 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | [11 empty] | 0
0
| 1 | 1 | 169 | 250 | 180 | 160 | [11 empty] | 1
| 1 | 2 | 171 | 250 | 160 | 170 | [11 empty] | 15
| 1 | 3 | 169 | 250 | 180 | 170 | [11 empty] | 16
| 1 | 4 | 171 | 250 | 180 | 170 | [11 empty] | 170.6667
| 1 | 5 | 172 | 250 | 180 | 170 | [11 empty] | 21
| 1 | 6 | 170 | 250 | 160 | 170
| Mean |  | 170.3333 | 250 | 173.3333 | 168.3333
| Median |  | 170.5 | 250 | 180 | 170
| Std Dev. |  | 1.2111 | 0 | 10.328 | 4.0825
| Rel.Std.Dev. |  | 0.711% | 0% | 5.958% | 2.425%
| PDM3 |  | -0.1953% | 46.48% | 1.562% | -1.367%
| Z-Score (Absolute) |  | 0.4 | 21.08 | 0.4 | 0.94
| Analytical results for Se in OREAS 465 (Indicative Value 33.6 ppm) | [16 empty] | Indicative
Se | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 1
1
| 1 | 1 | 26 | 40 | [13 empty] | 1
| 1 | 2 | 33 | 40 | [13 empty] | 25
| 1 | 3 | 28 | 40 | [13 empty] | 16
| 1 | 4 | 26 | 40 | [13 empty] | 33.5833
| 1 | 5 | 27 | 40 | [13 empty] | 22
| 1 | 6 | 23 | 40
| Mean |  | 27.1667 | 40
| Median |  | 26.5 | 40
| Std Dev. |  | 3.3116 | 0
| Rel.Std.Dev. |  | 12.19% | 0%
| PDM3 |  | -19.11% | 19.11%
| Z-Score (Absolute) |  | 0.67 | 0.67
| Analytical results for Si in OREAS 465 (Certified Value 1.53 wt.%) | [16 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 08 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | [7 empty] | wt.%
|  |  | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | [7 empty] | 2
3
| 1 | 1 | 1.9456 | 1.58 | 1.8605 | N.A. | 1.56 | 1.4886 | 1.608 | 1.56 | [7 empty] | 1
| 1 | 2 | 1.904 | 1.52 | 1.8652 | N.A. | 1.54 | 1.454 | 1.608 | 1.49 | [7 empty] | 9
| 1 | 3 | 1.9406 | 1.48 | 1.7948 | N.A. | 1.54 | 1.5096 | 1.594 | 1.51 | [7 empty] | 16
| 1 | 4 | 1.9258 | 1.51 | 1.8619 | N.A. | 1.54 | 1.4941 | 1.557 | 1.51 | [7 empty] | 1.5341
| 1 | 5 | 1.9594 | 1.48 | 1.8024 | N.A. | 1.56 | 1.4896 | 1.585 | 1.58 | [7 empty] | 46
| 1 | 6 | 1.9881 | 1.52 | 1.8705 | -- | 1.55 | 1.4706 | 1.594 | 1.54
| Mean |  | 1.9439 | 1.515 | 1.8425 |  | 1.5483 | 1.4844 | 1.591 | 1.5317
| Median |  | 1.9431 | 1.515 | 1.8612 |  | 1.545 | 1.4891 | 1.594 | 1.525
| Std Dev. |  | 0.02872 | 0.036742 | 0.034265 |  | 0.0098319 | 0.01941 | 0.01891 | 0.034303
| Rel.Std.Dev. |  | 1.477% | 2.425% | 1.86% |  | 0.635% | 1.308% | 1.189% | 2.24%
| PDM3 |  | 26.72% | -1.243% | 20.11% |  | 0.9301% | -3.238% | 3.711% | -0.1563%
| Z-Score (Absolute) |  | 6.26 | 0.53 | 4.66 | NA | 0 | 1.01 | 0.67 | 0.26
| Analytical results for Sm in OREAS 465 (Certified Value 1361 ppm) | [16 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 1360 | 1325 | 1400 | > 1000 | 1385 | 1354.8365 | 1351.7 | 1403 | 1188.5143 | 1290 | 1371 | 1294.7414 | 1080 | 1370 | 1360 | 1
| 1 | 2 | 1395 | 1310 | 1385 | > 1000 | 1410 | 1371.3786 | 1353.3 | 1390.1 | 1218.2138 | 1290 | 1351.5 | 1350.9999 | 1360 | 1300 | 1340 | 0
| 1 | 3 | 1395 | 1405 | 1400 | > 1000 | 1380 | 1382.9022 | 1332.7 | 1392.6 | 1186.3945 | 1290 | 1352.5 | 1351.5634 | 1370 | 1330 | 1390 | 16
| 1 | 4 | 1325 | 1370 | 1385 | > 1000 | 1355 | 1364.4342 | 1329.7 | 1386.3 | 1187.7805 | 1290 | 1350.5 | 1356.7045 | 1070 | 1470 | 1330 | 1361.0146
| 1 | 5 | 1400 | 1345 | 1400 | > 1000 | 1410 | 1368.2055 | 1323.6 | 1367.6 | 1201.0358 | 1290 | 1424 | 1348.7991 | 1390 | 1310 | 1320 | 47
| 1 | 6 | 1350 | 1380 | 1365 | > 1000 | 1410 | 1410.7459 | 1330.9 | 1415.8 | 1166.8524 | 1290 | 1336 | 1378.8016 | 1350 | 1400 | 1340
| Mean |  | 1370.8333 | 1355.8333 | 1389.1667 |  | 1391.6667 | 1375.4171 | 1336.9833 | 1392.5667 | 1191.4652 | 1290 | 1364.25 | 1346.935 | 1270 | 1363.3333 | 1346.6667
| Median |  | 1377.5 | 1357.5 | 1392.5 |  | 1397.5 | 1369.792 | 1331.8 | 1391.35 | 1188.1474 | 1290 | 1352 | 1351.2817 | 1355 | 1350 | 1340
| Std Dev. |  | 30.5641 | 35.6955 | 13.9344 |  | 22.5093 | 19.5789 | 12.4123 | 16.2317 | 17.0931 | 0 | 31.3157 | 27.8443 | 151.6575 | 64.3946 | 25.0333
| Rel.Std.Dev. |  | 2.23% | 2.633% | 1.003% |  | 1.617% | 1.423% | 0.9284% | 1.166% | 1.435% | 0% | 2.295% | 2.067% | 11.94% | 4.723% | 1.859%
| PDM3 |  | 0.7214% | -0.3807% | 2.068% |  | 2.252% | 1.058% | -1.766% | 2.318% | -12.46% | -5.218% | 0.2377% | -1.034% | -6.687% | 0.1704% | -1.054%
| Z-Score (Absolute) |  | 0.49 | 0.21 | 1.35 | NA | 1.47 | 0.71 | 1.1 | 1.51 | 7.92 | 3.13 | 0.38 | 0.14 | 0.34 | 0.14 | 0.64
| Analytical results for Sn in OREAS 465 (Certified Value 136 ppm) | [16 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*OES | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 153 | 143 | 156 | 149 | 144 | 152 | 109 | N.A. | 147 | 97.6098 | 69 | < 50 | 150 | 106 |  | 1
| 1 | 2 | 150 | 142 | 156 | 150 | 139 | 153 | 113 | N.A. | 144 | 95.8099 | 133 | < 50 | 140 | 109 |  | 10
| 1 | 3 | 158 | 155 | 148 | 150 | 143 | 150 | 113 | N.A. | 141 | 99.8381 | 136 | < 50 | 150 | 112 |  | 16
| 1 | 4 | 156 | 144 | 149 | 151 | 142 | 150 | 109 | N.A. | 148 | 107.4321 | 56 | < 50 | 160 | 111 |  | 135.8085
| 1 | 5 | 151 | 145 | 145 | 151 | 140 | 153 | 110 | N.A. | 151 | 108.7816 | 139 | < 50 | 150 | 106 |  | 48
| 1 | 6 | 153 | 149 | 147 | 148 | 142 | 150 | 115 | -- | 140 | 102.5418 | 131 | < 50 | 150 | 106
| Mean |  | 153.5 | 146.3333 | 150.1667 | 149.8333 | 141.6667 | 151.3333 | 111.5 |  | 145.1667 | 102.0022 | 110.6667 |  | 150 | 108.3333
| Median |  | 153 | 144.5 | 148.5 | 150 | 142 | 151 | 111.5 |  | 145.5 | 101.1899 | 132 |  | 150 | 107.5
| Std Dev. |  | 3.0166 | 4.8854 | 4.7081 | 1.169 | 1.8619 | 1.5055 | 2.51 |  | 4.2622 | 5.2546 | 37.6333 |  | 6.3246 | 2.7325
| Rel.Std.Dev. |  | 1.965% | 3.339% | 3.135% | 0.7802% | 1.314% | 0.9949% | 2.251% |  | 2.936% | 5.151% | 34.01% |  | 4.216% | 2.522%
| PDM3 |  | 13.03% | 7.75% | 10.57% | 10.33% | 4.314% | 11.43% | -17.9% |  | 6.891% | -24.89% | -18.51% |  | 10.45% | -20.23%
| Z-Score (Absolute) |  | 0.99 | 0.03 | 0.61 | 0.57 | 0.37 | 0.74 | 3.84 | NA | 0.03 | 4.93 | 2.68 | NA | 0.59 | 4.2
| Analytical results for Sr in OREAS 465 (Certified Value 5204 ppm) | [16 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
0
| 1 | 1 | 5470 | 5030 | 5320 | 5120 | 5280 | 5313 | 5035.4 | N.A. | 5400 | 6563.8475 | 4300 | 5280 | 5330 |  |  | 1
| 1 | 2 | 5320 | 4770 | 5310 | 5220 | 5080 | 5337 | 4995.7 | N.A. | 5300 | 7047.6758 | 4420 | 4920 | 5350 |  |  | 1
| 1 | 3 | 5530 | 5010 | 5060 | 5220 | 5270 | 5320 | 5024.2 | N.A. | 5300 | 6943.1463 | 4450 | 5040 | 5360 |  |  | 16
| 1 | 4 | 5630 | 5220 | 5010 | 5250 | 5060 | 5306 | 4987.4 | N.A. | 5300 | 7000.6203 | 4240 | 5240 | 5310 |  |  | 5203.8067
| 1 | 5 | 5460 | 4910 | 5040 | 5240 | 5070 | 5418 | 4982.1 | N.A. | 5300 | 7083.9713 | 4470 | 4990 | 5370 |  |  | 49
| 1 | 6 | 5560 | 4930 | 5020 | 5140 | 5170 | 5280 | 5029.6 | -- | 5300 | 7349.2162 | 4410 | 5010 | 5380
| Mean |  | 5495 | 4978.3333 | 5126.6667 | 5198.3333 | 5155 | 5329 | 5009.0667 |  | 5316.6667 | 6998.0796 | 4381.6667 | 5080 | 5350
| Median |  | 5500 | 4970 | 5050 | 5220 | 5125 | 5316.5 | 5009.95 |  | 5300 | 7024.148 | 4415 | 5025 | 5355
| Std Dev. |  | 105.9717 | 149.9889 | 146.924 | 54.5588 | 100.9455 | 47.4384 | 23.3212 |  | 40.8248 | 254.8394 | 91.086 | 145.4648 | 26.0768
| Rel.Std.Dev. |  | 1.929% | 3.013% | 2.866% | 1.05% | 1.958% | 0.8902% | 0.4656% |  | 0.7679% | 3.642% | 2.079% | 2.863% | 0.4874%
| PDM3 |  | 5.596% | -4.333% | -1.482% | -0.1052% | -0.9379% | 2.406% | -3.742% |  | 2.169% | 34.48% | -15.8% | -2.379% | 2.809%
| Z-Score (Absolute) |  | 1.34 | 0.84 | 0.21 | 0.09 | 0.09 | 0.64 | 0.71 | NA | 0.59 | 7.68 | 3.35 | 0.41 | 0.73
| Analytical results for Ta in OREAS 465 (Certified Value 79 ppm) | [16 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS |  |  | 0
1
| 1 | 1 | 76.1 | 79.2 | 81.8 | 81.9 | 81.7 | 73.8514 | 77.3 | N.A. | 24 | 52.4963 | 23 | 72.5 | 71.7 |  |  | 1
| 1 | 2 | 81 | 77.3 | 82.5 | 83.2 | 82.5 | 76.0757 | 76.4 | N.A. | 32 | 50.7522 | 89.5 | 73.5 | 72.4 |  |  | 27
| 1 | 3 | 78.2 | 78.6 | 81.8 | 83.8 | 79.8 | 75.4958 | 72.1 | N.A. | 23 | 53.8047 | 90.5 | 73.5 | 74.4 |  |  | 16
| 1 | 4 | 74.8 | 82.5 | 81.2 | 82.9 | 79.8 | 73.8435 | 75.6 | N.A. | 24 | 56.5426 | 19.2 | 72.5 | 71.3 |  |  | 78.6692
| 1 | 5 | 77.6 | 76.2 | 81.5 | 82.6 | 84.2 | 77.6924 | 74.9 | N.A. | 31 | 57.2855 | 91.1 | 71.5 | 71.1 |  |  | 50
| 1 | 6 | 74.4 | 76.1 | 80.3 | 81.2 | 84.8 | 74.9929 | 79.4 | -- | 24 | 55.1738 | 86.5 | 70.5 | 71.7
| Mean |  | 77.0167 | 78.3167 | 81.5167 | 82.6 | 82.1333 | 75.3253 | 75.95 |  | 26.3333 | 54.3425 | 66.6333 | 72.3333 | 72.1
| Median |  | 76.85 | 77.95 | 81.65 | 82.75 | 82.1 | 75.2443 | 76 |  | 24 | 54.4893 | 88 | 72.5 | 71.7
| Std Dev. |  | 2.458 | 2.3995 | 0.73598 | 0.93167 | 2.1257 | 1.4611 | 2.4485 |  | 4.0332 | 2.4812 | 35.3259 | 1.169 | 1.2116
| Rel.Std.Dev. |  | 3.191% | 3.064% | 0.9029% | 1.128% | 2.588% | 1.94% | 3.224% |  | 15.32% | 4.566% | 53.02% | 1.616% | 1.68%
| PDM3 |  | -2.101% | -0.4481% | 3.62% | 4.997% | 4.403% | -4.251% | -3.456% |  | -66.53% | -30.92% | -15.3% | -8.054% | -8.35%
| Z-Score (Absolute) |  | 0.08 | 0.26 | 0.72 | 0.88 | 0.81 | 0.17 | 0.08 | NA | 7.18 | 3.17 | 1.85 | 0.59 | 0.63
| Analytical results for Tb in OREAS 465 (Certified Value 57 ppm) | [16 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 0
1
| 1 | 1 | 52.7 | 56.2 | 58.5 | 56.8 | 59 | 59.9366 | 55.9 | 57.5 | 50.0417 | 60 | 57.8 | 66.9813 | 43.3 | 60 | 55.9 | 1
| 1 | 2 | 55.1 | 54.6 | 60.8 | 57.1 | 54.7 | 60.7599 | 55.3 | 57.2 | 51.0982 | 60 | 58.2 | 71.5391 | 54.9 | 58 | 54.9 | 3
| 1 | 3 | 54.7 | 57.8 | 57.9 | 57.3 | 58.3 | 59.7822 | 54.6 | 55.4 | 50.7889 | 60 | 57.1 | 69.4945 | 54.9 | 59.2 | 58 | 16
| 1 | 4 | 51.2 | 54.8 | 69.8 | 57.4 | 55.7 | 58.7446 | 55.2 | 56.3 | 50.5896 | 60 | 56.7 | 69.8335 | 42.4 | 60.4 | 52.9 | 56.5714
| 1 | 5 | 54 | 56.4 | 68.3 | 57.1 | 55 | 59.2075 | 54.3 | 56.5 | 50.4406 | 60 | 61.6 | 70.2535 | 54.8 | 58.6 | 53.5 | 51
| 1 | 6 | 51.7 | 55.6 | 70.4 | 55.4 | 57.7 | 61.3588 | 54 | 57.1 | 49.6703 | 60 | 56.7 | 71.7617 | 54.1 | 61.2 | 54.2
| Mean |  | 53.2333 | 55.9 | 64.2833 | 56.85 | 56.7333 | 59.9649 | 54.8833 | 56.6667 | 50.4382 | 60 | 58.0167 | 69.9773 | 50.7333 | 59.5667 | 54.9
| Median |  | 53.35 | 55.9 | 64.55 | 57.1 | 56.7 | 59.8594 | 54.9 | 56.8 | 50.5151 | 60 | 57.45 | 70.0435 | 54.45 | 59.6 | 54.55
| Std Dev. |  | 1.612 | 1.1781 | 5.8362 | 0.73959 | 1.8294 | 0.96665 | 0.70828 | 0.76594 | 0.51545 | 0 | 1.8563 | 1.7279 | 6.1203 | 1.1894 | 1.8472
| Rel.Std.Dev. |  | 3.028% | 2.108% | 9.079% | 1.301% | 3.225% | 1.612% | 1.291% | 1.352% | 1.022% | 0% | 3.2% | 2.469% | 12.06% | 1.997% | 3.365%
| PDM3 |  | -5.901% | -1.187% | 13.63% | 0.4924% | 0.2862% | 5.999% | -2.984% | 0.1683% | -10.84% | 6.061% | 2.555% | 23.7% | -10.32% | 5.295% | -2.955%
| Z-Score (Absolute) |  | 1.15 | 0.27 | 1.52 | 0.04 | 0 | 1.06 | 0.61 | 0.02 | 2.06 | 0.91 | 0.19 | 4.34 | 0.67 | 0.93 | 0.6
| Analytical results for Te in OREAS 465 (Indicative Value < 1 ppm) | [16 empty] | Indicative
Te | Round | Replicate | Lab | Lab | [13 empty] | 1
| No | No | 09 | 22 | [13 empty] | ppm
|  |  | PF*MS | PF*MS | [13 empty] | 2
2
| 1 | 1 | < 1 | < 1 | [13 empty] | 1
| 1 | 2 | < 1 | < 1 | [13 empty] | 17
| 1 | 3 | < 1 | < 1 | [13 empty] | 16
| 1 | 4 | < 1 | < 1 | [13 empty] | < 1
| 1 | 5 | < 1 | < 1 | [13 empty] | 23
| 1 | 6 | < 1 | < 1
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA
| Analytical results for Th in OREAS 465 (Certified Value 866 ppm) | [16 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 0
0
| 1 | 1 | 886 | 889 | 898 | 870 | 920 | 940.5207 | 890.7 | 854.1 | 607.7705 | 814 | 846.5 | 733.5394 | 725 | 812 | 784 | 1
| 1 | 2 | 877 | 819 | 904 | 876 | 859 | 952.4315 | 889.3 | 858.2 | 607.0136 | 847 | 851 | 739.8398 | 705 | 777 | 783 | 0
| 1 | 3 | 865 | 857 | 907 | 876 | 914 | 960.5837 | 884.6 | 834 | 605.7115 | 853 | 838 | 740.4677 | 704 | 834 | 810 | 16
| 1 | 4 | 860 | 914 | 920 | 888 | 896 | 901.8899 | 884.5 | 837.6 | 605.7779 | 859 | 854 | 753.5548 | 718 | 877 | 774 | 865.5289
| 1 | 5 | 864 | 830 | 885 | 868 | 898 | 949.986 | 879.1 | 829.7 | 606.9351 | 859 | 876 | 762.1024 | 724 | 813 | 784 | 52
| 1 | 6 | 884 | 827 | 899 | 849 | 908 | 928.5664 | 874.8 | 849 | 602.4423 | 857 | 825 | 757.9223 | 793 | 878 | 765
| Mean |  | 872.6667 | 856 | 902.1667 | 871.1667 | 899.1667 | 938.9964 | 883.8333 | 843.7667 | 605.9418 | 848.1667 | 848.4167 | 747.9044 | 728.1667 | 831.8333 | 783.3333
| Median |  | 871 | 843.5 | 901.5 | 873 | 903 | 945.2533 | 884.55 | 843.3 | 606.3565 | 855 | 848.75 | 747.0113 | 721 | 823.5 | 783.5
| Std Dev. |  | 11.1295 | 38.3145 | 11.5484 | 12.9061 | 21.711 | 21.2381 | 6.0345 | 11.6084 | 1.8871 | 17.3484 | 17.0834 | 11.494 | 33.0298 | 39.8267 | 15.0687
| Rel.Std.Dev. |  | 1.275% | 4.476% | 1.28% | 1.481% | 2.415% | 2.262% | 0.6828% | 1.376% | 0.3114% | 2.045% | 2.014% | 1.537% | 4.536% | 4.788% | 1.924%
| PDM3 |  | 0.8247% | -1.101% | 4.233% | 0.6514% | 3.886% | 8.488% | 2.115% | -2.514% | -29.99% | -2.006% | -1.977% | -13.59% | -15.87% | -3.893% | -9.497%
| Z-Score (Absolute) |  | 0.41 | 0.02 | 1.1 | 0.38 | 1.03 | 1.96 | 0.67 | 0.26 | 5.82 | 0 | 0.15 | 2.5 | 3.27 | 0.54 | 1.68
| Analytical results for Ti in OREAS 465 (Certified Value 6.3 wt.%) | [16 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 21 | 22 | 23 |  |  |  |  | wt.%
|  |  | PF*OES | PF*OES | BF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 6.487 | 6.24 | 6.17 | 5.9358 | N.A. | 6.69 | 6.1809 | 1.36 | 6.461 | 6.46 | 6.33 |  |  |  |  | 1
| 1 | 2 | 6.554 | 6.2 | 6.18 | 5.8756 | N.A. | 6.58 | 6.0428 | 5.79 | 6.391 | 6.31 | 6.36 |  |  |  |  | 4
| 1 | 3 | 6.487 | 6.22 | 6.2 | 6.119 | N.A. | 6.52 | 6.2039 | 5.88 | 6.323 | 6.44 | 6.35 |  |  |  |  | 16
| 1 | 4 | 6.399 | 6.26 | 6.23 | 5.928 | N.A. | 6.61 | 6.2212 | 1.13 | 6.323 | 6.42 | 6.4 |  |  |  |  | 6.3017
| 1 | 5 | 6.388 | 6.25 | 6.3 | 5.9803 | N.A. | 6.58 | 6.2442 | 5.83 | 6.405 | 6.33 | 6.38 |  |  |  |  | 53
| 1 | 6 | 6.192 | 6.23 | 6.23 | 5.8879 | -- | 6.58 | 6.1809 | 5.56 | 6.38 | 6.26 | 6.39
| Mean |  | 6.4178 | 6.2333 | 6.2183 | 5.9544 |  | 6.5933 | 6.179 | 4.2583 | 6.3805 | 6.37 | 6.3683
| Median |  | 6.443 | 6.235 | 6.215 | 5.9319 |  | 6.58 | 6.1924 | 5.675 | 6.3855 | 6.375 | 6.37
| Std Dev. |  | 0.12677 | 0.021602 | 0.047081 | 0.088825 |  | 0.055737 | 0.071015 | 2.3378 | 0.052542 | 0.080994 | 0.026394
| Rel.Std.Dev. |  | 1.975% | 0.3466% | 0.7571% | 1.492% |  | 0.8454% | 1.149% | 54.9% | 0.8235% | 1.271% | 0.4145%
| PDM3 |  | 1.843% | -1.085% | -1.323% | -5.51% |  | 4.628% | -1.947% | -32.43% | 1.251% | 1.084% | 1.058%
| Z-Score (Absolute) |  | 0.79 | 0.46 | 0.56 | 2.34 | NA | 1.98 | 0.82 | 13.81 | 0.54 | 0.47 | 0.46
| Analytical results for Tl in OREAS 465 (Indicative Value < 0.5 ppm) | [16 empty] | Indicative
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 09 | 14 | 17 | 20 | 22 | 23 | [9 empty] | ppm
|  |  | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | N.A. | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 1
| 1 | 2 | < 0.5 | N.A. | 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 18
| 1 | 3 | < 0.5 | N.A. | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 16
| 1 | 4 | < 0.5 | N.A. | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | < 0.5
| 1 | 5 | < 0.5 | N.A. | < 0.1 | < 0.5 | < 2 | < 0.5 | [9 empty] | 24
| 1 | 6 | < 0.5 | -- | < 0.1 | < 0.5 | < 2 | < 0.5
| Mean |  |  |  | 0.1
| Median |  |  |  | 0.1
Std Dev.
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA
| Analytical results for Tm in OREAS 465 (Certified Value 4.52 ppm) | [16 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 2
3
| 1 | 1 | 4.38 | 4.37 | 4.91 | 4.73 | 4.48 | 4.3422 | 4.3 | 4.9 | 3.0616 | < 44 | 4.6 | 3.453 | 3.92 | 4.6 | 4.27 | 1
| 1 | 2 | 4.5 | 4.22 | 5.02 | 4.85 | 4.49 | 4.4304 | 4.4 | 4.8 | 3.0894 | < 44 | 4.5 | 3.5056 | 4.71 | 4 | 4.63 | 5
| 1 | 3 | 4.4 | 4.35 | 4.78 | 4.68 | 4.37 | 4.0567 | 4.5 | 4.9 | 2.9974 | < 44 | 4.4 | 3.5283 | 4.71 | 4.4 | 4.26 | 16
| 1 | 4 | 4.21 | 4.59 | 5.01 | 4.93 | 4.27 | 4.3596 | 4.6 | 5.4 | 2.9756 | < 44 | 4.4 | 3.5948 | 3.79 | 5.4 | 3.8 | 4.5221
| 1 | 5 | 4.32 | 4.19 | 4.73 | 4.67 | 4.65 | 4.4315 | 4.4 | 4.8 | 2.4671 | < 44 | 4.6 | 3.5509 | 4.81 | 4.2 | 4.31 | 54
| 1 | 6 | 4.18 | 4.32 | 4.64 | 4.7 | 4.64 | 4.3394 | 4.5 | 4.9 | 2.9141 | < 44 | 4.4 | 3.613 | 4.73 | 5 | 4.6
| Mean |  | 4.3317 | 4.34 | 4.8483 | 4.76 | 4.4833 | 4.3266 | 4.45 | 4.95 | 2.9175 |  | 4.4833 | 3.5409 | 4.445 | 4.6 | 4.3117
| Median |  | 4.35 | 4.335 | 4.845 | 4.715 | 4.485 | 4.3509 | 4.45 | 4.9 | 2.9865 |  | 4.45 | 3.5396 | 4.71 | 4.5 | 4.29
| Std Dev. |  | 0.12106 | 0.14199 | 0.15587 | 0.10583 | 0.14882 | 0.13863 | 0.10488 | 0.22583 | 0.22931 |  | 0.098319 | 0.058879 | 0.46034 | 0.52154 | 0.30023
| Rel.Std.Dev. |  | 2.795% | 3.272% | 3.215% | 2.223% | 3.319% | 3.204% | 2.357% | 4.562% | 7.86% |  | 2.193% | 1.663% | 10.36% | 11.34% | 6.963%
| PDM3 |  | -4.211% | -4.026% | 7.215% | 5.261% | -0.8568% | -4.322% | -1.594% | 9.463% | -35.48% |  | -0.8568% | -21.7% | -1.704% | 1.723% | -4.653%
| Z-Score (Absolute) |  | 0.63 | 0.58 | 2.23 | 1.74 | 0.21 | 0.36 | 0.03 | 2.29 | 7.93 | NA | 0.21 | 4.99 | 0.72 | 0.03 | 0.74
| Analytical results for U in OREAS 465 (Certified Value 13.6 ppm) | [16 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 13.05 | 13.4 | 13.5 | 13.8 | 13.6 | 13.7447 | 13.2 | 13.8 | 11.5243 | 13 | 14 | 12.9753 | 5.82 | 12.5 | 14 | 1
| 1 | 2 | 12.8 | 13.5 | 13.9 | 13.85 | 13.6 | 13.9351 | 13.4 | 13.8 | 11.4201 | 13 | 13.5 | 13.5801 | 8.77 | 12 | 13.3 | 6
| 1 | 3 | 13.5 | 13.9 | 13.2 | 14.05 | 13.3 | 14.2437 | 13.4 | 13.5 | 11.7594 | 13 | 13.5 | 13.4656 | 8.88 | 12.5 | 14.2 | 16
| 1 | 4 | 13.45 | 13.3 | 14.1 | 14.05 | 13.3 | 14.2156 | 13.6 | 13.6 | 11.7796 | 13 | 13.5 | 13.5245 | 5.63 | 13 | 13.1 | 13.6006
| 1 | 5 | 13.35 | 13.7 | 14.3 | 13.7 | 13.5 | 14.5807 | 13.3 | 13.6 | 12.0074 | 13 | 14.5 | 13.5369 | 9.11 | 12 | 13.9 | 55
| 1 | 6 | 13.45 | 13.4 | 13.8 | 13.35 | 14.2 | 14.4386 | 13.4 | 14 | 11.5709 | 13 | 13 | 13.7108 | 8.54 | 13 | 13.7
| Mean |  | 13.2667 | 13.5333 | 13.8 | 13.8 | 13.5833 | 14.1931 | 13.3833 | 13.7167 | 11.677 | 13 | 13.6667 | 13.4655 | 7.7917 | 12.5 | 13.7
| Median |  | 13.4 | 13.45 | 13.85 | 13.825 | 13.55 | 14.2297 | 13.4 | 13.7 | 11.6652 | 13 | 13.5 | 13.5307 | 8.655 | 12.5 | 13.8
| Std Dev. |  | 0.28048 | 0.22509 | 0.4 | 0.26077 | 0.33116 | 0.31023 | 0.13292 | 0.18348 | 0.21311 | 1.9459e-15 | 0.5164 | 0.25384 | 1.6124 | 0.44721 | 0.42426
| Rel.Std.Dev. |  | 2.114% | 1.663% | 2.899% | 1.89% | 2.438% | 2.186% | 0.9931% | 1.338% | 1.825% | 1.497e-14% | 3.779% | 1.885% | 20.69% | 3.578% | 3.097%
| PDM3 |  | -2.455% | -0.4942% | 1.466% | 1.466% | -0.1266% | 4.357% | -1.597% | 0.8537% | -14.14% | -4.416% | 0.4861% | -0.9928% | -42.71% | -8.092% | 0.7312%
| Z-Score (Absolute) |  | 0.85 | 0.09 | 0.67 | 0.67 | 0.06 | 1.8 | 0.51 | 0.44 | 5.38 | 1.61 | 0.29 | 0 | 16.46 | 3.03 | 0.39
| Analytical results for V in OREAS 465 (Certified Value 534 ppm) | [16 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES | BF*OES | PF*OES | PF*OES/MS |  |  | 0
0
| 1 | 1 | 495 | 530 | 575 | 568 | 567 | 537 | 557 | N.A. | 600 | 514.2 | 563 | 500 | 489 |  |  | 1
| 1 | 2 | 486 | 502 | 579 | 576 | 556 | 545 | 536 | N.A. | 600 | 495.7 | 519 | 500 | 487 |  |  | 7
| 1 | 3 | 501 | 521 | 566 | 589 | 562 | 547 | 532 | N.A. | 600 | 524.9 | 512 | 500 | 501 |  |  | 16
| 1 | 4 | 491 | 531 | 555 | 589 | 529 | 548 | 541 | N.A. | 600 | 513.1 | 542 | 500 | 480 |  |  | 534.1117
| 1 | 5 | 488 | 513 | 548 | 589 | 527 | 547 | 573 | N.A. | 600 | 512 | 512 | 500 | 484 |  |  | 56
| 1 | 6 | 486 | 529 | 550 | 576 | 549 | 544 | 539 | -- | 600 | 507.5 | 518 | 500 | 487
| Mean |  | 491.1667 | 521 | 562.1667 | 581.1667 | 548.3333 | 544.6667 | 546.3333 |  | 600 | 511.2333 | 527.6667 | 500 | 488
| Median |  | 489.5 | 525 | 560.5 | 582.5 | 552.5 | 546 | 540 |  | 600 | 512.55 | 518.5 | 500 | 487
| Std Dev. |  | 5.9133 | 11.5758 | 13.1365 | 9.0646 | 16.8721 | 4.0332 | 15.6162 |  | 0 | 9.5381 | 20.5589 | 6.2269e-14 | 7.0993
| Rel.Std.Dev. |  | 1.204% | 2.222% | 2.337% | 1.56% | 3.077% | 0.7405% | 2.858% |  | 0% | 1.866% | 3.896% | 1.245e-14% | 1.455%
| PDM3 |  | -8.04% | -2.455% | 5.253% | 8.81% | 2.663% | 1.976% | 2.288% |  | 12.34% | -4.283% | -1.207% | -6.387% | -8.633%
| Z-Score (Absolute) |  | 1.06 | 0.27 | 0.83 | 1.33 | 0.46 | 0.36 | 0.27 | NA | 1.83 | 0.52 | 0.28 | 0.82 | 1.14
| Analytical results for W in OREAS 465 (Certified Value 7.52 ppm) | [16 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 17 | 20 | 22 | 23 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*MS | PF*MS | PF*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 8 | 6 | 8 | 7 | 8 | 7 | 8 | 10 | 3 | 5 | 10 |  |  |  |  | 1
| 1 | 2 | 13 | 7 | 7 | 7 | 7 | 7 | 8 | < 10 | 2 | 5 | 9 |  |  |  |  | 8
| 1 | 3 | 7 | 7 | 9 | 7 | 7 | 7 | 8 | < 10 | 3 | 5 | 9 |  |  |  |  | 16
| 1 | 4 | 7 | 7 | 8 | 7 | 7 | 7 | 8 | < 10 | < 1 | 5 | 10 |  |  |  |  | 7.52
| 1 | 5 | 7 | 6 | 8 | 7 | 7 | 7 | 8 | < 10 | < 1 | 5 | 9 |  |  |  |  | 57
| 1 | 6 | 7 | 6 | 8 | 7 | 7 | 7 | 8 | < 10 | 2 | 5 | 9
| Mean |  | 8.1667 | 6.5 | 8 | 7 | 7.1667 | 7 | 8 | 10 | 2.5 | 5 | 9.3333
| Median |  | 7 | 6.5 | 8 | 7 | 7 | 7 | 8 | 10 | 2.5 | 5 | 9
| Std Dev. |  | 2.4014 | 0.54772 | 0.63246 | 9.7295e-16 | 0.40825 | 9.7295e-16 | 1.9459e-15 |  | 0.57735 | 0 | 0.5164
| Rel.Std.Dev. |  | 29.4% | 8.427% | 7.906% | 1.39e-14% | 5.696% | 1.39e-14% | 2.432e-14% |  | 23.09% | 0% | 5.533%
| PDM3 |  | 8.599% | -13.56% | 6.383% | -6.915% | -4.699% | -6.915% | 6.383% | 32.98% | -66.76% | -33.51% | 24.11%
| Z-Score (Absolute) |  | 0.03 | 0.54 | 0.67 | 0.13 | 0 | 0.13 | 0.67 | 2.29 | 3.78 | 1.75 | 1.75
| Analytical results for Y in OREAS 465 (Certified Value 524 ppm) | [16 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 14 | 17 | 18 | 20 | 22 | 23 |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | BF*MS | PF*OES | BF*OES | BF*MS | PF*MS | PF*OES/MS |  | 0
0
| 1 | 1 | 530 | 559 | 568 | 539 | 534 | 557.8149 | 538.9 | 509.7 | 504 | 511 | 457.5 | 155 | 577 | 482 |  | 1
| 1 | 2 | 534 | 507 | 571 | 547 | 546 | 544.8898 | 537.8 | 517.4 | 512 | 502.5 | 446.5 | 468 | 548 | 474 |  | 0
| 1 | 3 | 515 | 549 | 541 | 549 | 544 | 578.0044 | 537.1 | 508.2 | 504 | 498.5 | 462.2 | 49 | 565 | 489 |  | 16
| 1 | 4 | 502 | 533 | 537 | 552 | 535 | 552.8914 | 543.9 | 520.6 | 512 | 497.5 | 459.1 | 56.7 | 584 | 470 |  | 523.8165
| 1 | 5 | 523 | 560 | 541 | 549 | 548 | 545.8984 | 543.6 | 522.5 | 496 | 528 | 453.6 | 319 | 563 | 473 |  | 58
| 1 | 6 | 510 | 538 | 537 | 543 | 558 | 551.5735 | 537.2 | 535.8 | 496 | 494 | 444.1 | 322 | 569 | 472
| Mean |  | 519 | 541 | 549.1667 | 546.5 | 544.1667 | 555.1787 | 539.75 | 519.0333 | 504 | 505.25 | 453.8333 | 228.2833 | 567.6667 | 476.6667
| Median |  | 519 | 543.5 | 541 | 548 | 545 | 552.2325 | 538.35 | 519 | 504 | 500.5 | 455.55 | 237 | 567 | 473.5
| Std Dev. |  | 12.2311 | 19.8897 | 15.8798 | 4.7223 | 8.9088 | 12.1505 | 3.1653 | 10.0273 | 7.1554 | 12.5728 | 7.2077 | 168.1845 | 12.4204 | 7.3121
| Rel.Std.Dev. |  | 2.357% | 3.676% | 2.892% | 0.8641% | 1.637% | 2.189% | 0.5864% | 1.932% | 1.42% | 2.488% | 1.588% | 73.67% | 2.188% | 1.534%
| PDM3 |  | -0.9195% | 3.28% | 4.84% | 4.33% | 3.885% | 5.987% | 3.042% | -0.9131% | -3.783% | -3.544% | -13.36% | -56.42% | 8.371% | -9.001%
| Z-Score (Absolute) |  | 0.34 | 0.38 | 0.65 | 0.56 | 0.49 | 0.7 | 0.34 | 0.34 | 0.84 | 0.94 | 2.49 | 9.91 | 1.26 | 1.82
| Analytical results for Yb in OREAS 465 (Certified Value 19 ppm) | [16 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 11 | 12 | 14 | 17 | 18 | 20 | 22 | 23 | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | PF*MS | BF*MS | PF*MS | BF*MS | PF*OES | BF*OES/MS | BF*MS | PF*MS | PF*OES/MS | 1
2
| 1 | 1 | 18.8 | 19 | 20.7 | 19.45 | 18.4 | 17.6473 | 19.1 | 19.6 | 18.1088 | < 90 | 20.4 | 17.7327 | 5.5 | 19.5 | 21.9 | 1
| 1 | 2 | 18.45 | 18.2 | 18.8 | 18.9 | 19.2 | 18.8413 | 19.4 | 19.4 | 18.4281 | < 90 | 19.8 | 18.2979 | 13.1 | 18 | 21.8 | 9
| 1 | 3 | 18.2 | 19.6 | 19.4 | 19.1 | 18 | 19.3305 | 18.6 | 19.3 | 17.774 | < 90 | 19.7 | 18.2388 | 1.4 | 19 | 23.8 | 16
| 1 | 4 | 18.6 | 18.8 | 20.2 | 19.9 | 18.5 | 18.3735 | 19.1 | 21.9 | 17.958 | < 90 | 20.1 | 18.2423 | 1.5 | 20 | 22.7 | 18.9818
| 1 | 5 | 18.35 | 19 | 18.5 | 19.3 | 18.9 | 18.8223 | 18.3 | 20 | 15.1346 | < 90 | 21.9 | 18.2422 | 14.8 | 18 | 22 | 59
| 1 | 6 | 18.9 | 19 | 20.3 | 19.5 | 19.4 | 19.3496 | 19.2 | 20.3 | 18.0825 | < 90 | 19.8 | 18.5166 | 14 | 19 | 22.7
| Mean |  | 18.55 | 18.9333 | 19.65 | 19.3583 | 18.7333 | 18.7274 | 18.95 | 20.0833 | 17.581 |  | 20.2833 | 18.2117 | 8.3833 | 18.9167 | 22.4833
| Median |  | 18.525 | 19 | 19.8 | 19.375 | 18.7 | 18.8318 | 19.1 | 19.8 | 18.0203 |  | 19.95 | 18.2423 | 9.3 | 19 | 22.35
| Std Dev. |  | 0.26833 | 0.45019 | 0.88713 | 0.34701 | 0.52789 | 0.64258 | 0.41352 | 0.96626 | 1.2175 |  | 0.83287 | 0.25786 | 6.3155 | 0.80104 | 0.75741
| Rel.Std.Dev. |  | 1.447% | 2.378% | 4.515% | 1.793% | 2.818% | 3.431% | 2.182% | 4.811% | 6.925% |  | 4.106% | 1.416% | 75.33% | 4.235% | 3.369%
| PDM3 |  | -2.275% | -0.2551% | 3.52% | 1.984% | -1.309% | -1.34% | -0.1673% | 5.803% | -7.379% |  | 6.857% | -4.057% | -55.83% | -0.3429% | 18.45%
| Z-Score (Absolute) |  | 0.44 | 0.01 | 0.85 | 0.51 | 0.23 | 0.23 | 0.03 | 0.94 | 1.01 | NA | 1.22 | 0.84 | 12.4 | 0.01 | 4.19
| Analytical results for Zn in OREAS 465 (Indicative Value 1097 ppm) | [16 empty] | Indicative
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 14 | 17 | 18 | 21 | 22 | 23 | [9 empty] | ppm
|  |  | BF*OES | PF*OES | BF*OES/MS | PF*OES | PF*OES | PF*OES/MS | [9 empty] | 0
0
| 1 | 1 | N.A. | 1100 | 1536.2769 | 1150 | 1200 | 1030 | [9 empty] | 1
| 1 | 2 | N.A. | 1100 | 1652.6865 | 1110 | 1150 | 1020 | [9 empty] | 10
| 1 | 3 | N.A. | 1100 | 1607.5296 | 1160 | 1150 | 1020 | [9 empty] | 16
| 1 | 4 | N.A. | 1100 | 1716.1828 | 1110 | 1150 | 1010 | [9 empty] | 1097.0833
| 1 | 5 | N.A. | 1100 | 1665.5466 | 1120 | 1100 | 1020 | [9 empty] | 25
| 1 | 6 | -- | 1100 | 1565.3534 | 1100 | 1100 | 1030
| Mean |  |  | 1100 | 1623.9293 | 1125 | 1141.6667 | 1021.6667
| Median |  |  | 1100 | 1630.108 | 1115 | 1150 | 1020
| Std Dev. |  |  | 2.4908e-13 | 67.0115 | 24.2899 | 37.6386 | 7.5277
| Rel.Std.Dev. |  |  | 2.264e-14% | 4.127% | 2.159% | 3.297% | 0.7368%
| PDM3 |  |  | 0.2659% | 48.02% | 2.545% | 4.064% | -6.874%
| Z-Score (Absolute) |  | NA | 0.67 | 13.46 | 0 | 0.45 | 2.79
| Analytical results for Zr in OREAS 465 (Certified Value 1880 ppm) | [16 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 09 | 11 | 14 | 17 | 18 | 20 |  |  |  |  | ppm
|  |  | BF*MS | BF*MS | BF*MS | BF*MS | BF*MS | PF*MS | BF*MS | BF*OES | PF*OES | BF*OES | BF*MS |  |  |  |  | 0
0
| 1 | 1 | 1970 | 1980 | 2040 | 1985 | 2010 | 810 | 1397 | N.A. | 1183 | 1821.9 | 662 |  |  |  |  | 1
| 1 | 2 | 2040 | 1870 | 2080 | 2050 | 2000 | 832 | 1536 | N.A. | 1209 | 1650.1 | 494 |  |  |  |  | 11
| 1 | 3 | 1950 | 1930 | 1960 | 2040 | 2000 | 932 | 1549 | N.A. | 1352 | 1231.1 | 616 |  |  |  |  | 16
| 1 | 4 | 1910 | 1980 | 2000 | 2070 | 1960 | 819 | 1347 | N.A. | 1357 | 1776 | 140 |  |  |  |  | 1879.6006
| 1 | 5 | 1960 | 1880 | 1985 | 2070 | 1960 | 950 | 1374 | N.A. | 1180 | 1849.3 | 836 |  |  |  |  | 60
| 1 | 6 | 1910 | 1910 | 1945 | 2020 | 2090 | 890 | 1500 | -- | 1308 | 1807.1 | 1260
| Mean |  | 1956.6667 | 1925 | 2001.6667 | 2039.1667 | 2003.3333 | 872.1667 | 1450.5 |  | 1264.8333 | 1689.25 | 668
| Median |  | 1955 | 1920 | 1992.5 | 2045 | 2000 | 861 | 1448.5 |  | 1258.5 | 1791.55 | 639
| Std Dev. |  | 48.0278 | 47.6445 | 50.6623 | 32.6216 | 47.6095 | 60.4332 | 88.1924 |  | 83.6264 | 234.9725 | 371.6149
| Rel.Std.Dev. |  | 2.455% | 2.475% | 2.531% | 1.6% | 2.377% | 6.929% | 6.08% |  | 6.612% | 13.91% | 55.63%
| PDM3 |  | 4.1% | 2.415% | 6.494% | 8.489% | 6.583% | -53.6% | -22.83% |  | -32.71% | -10.13% | -64.46%
| Z-Score (Absolute) |  | 0.42 | 0.29 | 0.6 | 0.75 | 0.6 | 3.93 | 1.61 | NA | 2.36 | 0.29 | 4.75
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 465 (Certified Value 5.48 ppm) | [15 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 5.42 | 5.68 | 5.3 | 5.71 | 5.36 | 5.4899 | 5.6 | 8.4 | < 0.01 | 6.92 | 5 | 5 |  |  | 1
| 1 | 2 | 5.72 | 5.23 | 5.09 | 5.59 | 5.22 | 5.6133 | 5.7 | 5.8 | < 0.01 | 7.74 | 5 | 5 |  |  | 15
| 1 | 3 | 5.54 | 5.49 | 5.31 | 6.02 | 5.09 | 5.7798 | 5.2 | 5.7 | < 0.01 | 7.53 | 5 | 5 |  |  | 16
| 1 | 4 | 5.44 | 5.7 | 5.52 | 5.94 | 5.01 | 5.6111 | 5.5 | 6.1 | < 0.01 | 6.5 | 5 | 6 |  |  | 5.4832
| 1 | 5 | 5.68 | 5.79 | 5.59 | 5.47 | 5.25 | 5.597 | 5 | 6.3 | < 0.01 | 6.12 | 5 | 5.5 |  |  | 62
| 1 | 6 | 6.26 | 5.78 | 5.05 | 5.63 | 5.18 | 5.7192 | 5.1 | 5.7 | < 0.01 | 6.74 | 6 | 5
| Mean |  | 5.6767 | 5.6117 | 5.31 | 5.7267 | 5.185 | 5.6351 | 5.35 | 6.3333 |  | 6.925 | 5.1667 | 5.25
| Median |  | 5.61 | 5.69 | 5.305 | 5.67 | 5.2 | 5.6122 | 5.35 | 5.95 |  | 6.83 | 5 | 5
| Std Dev. |  | 0.31072 | 0.21591 | 0.21845 | 0.21248 | 0.12309 | 0.10157 | 0.2881 | 1.0405 |  | 0.61533 | 0.40825 | 0.41833
| Rel.Std.Dev. |  | 5.474% | 3.848% | 4.114% | 3.71% | 2.374% | 1.802% | 5.385% | 16.43% |  | 8.886% | 7.902% | 7.968%
| PDM3 |  | 3.529% | 2.343% | -3.158% | 4.441% | -5.438% | 2.77% | -2.429% | 15.5% |  | 26.3% | -5.772% | -4.253%
| Z-Score (Absolute) |  | 0.15 | 0 | 0.67 | 0.26 | 0.95 | 0.05 | 0.58 | 0.69 | NA | 2.94 | 0.99 | 0.81
| Analytical results for Al in OREAS 465 (Certified Value 6.21 wt.%) | [15 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 5.83 | 5.99 | 6.23 | 6.82 | 6.06 | 6.4166 | 6.48 | 5.34 | 6.73 | 6.1773 | 5.8912 | 6.26 | 5.608 | 6.94 | 1
| 1 | 2 | 5.97 | 5.8 | 6.07 | 6.55 | 6.24 | 6.3413 | 6.55 | 5.07 | 6.66 | 5.9805 | 5.9703 | 6.12 | 6.054 | 6.8 | 1
| 1 | 3 | 4.9 | 5.96 | 6.08 | 6.73 | 6.18 | 6.3775 | 6.52 | 5.36 | 6.72 | 6.0338 | 6.0144 | 5.87 | 6.009 | 6.84 | 16
| 1 | 4 | 5.98 | 6.03 | 6.45 | 6.67 | 6.26 | 6.3069 | 6.45 | 6.38 | 6.77 | 6.1946 | 5.9817 | 6.08 | 5.804 | 6.85 | 6.2092
| 1 | 5 | 5.19 | 6.01 | 6.35 | 6.6 | 6.12 | 6.246 | 6.52 | 6.34 | 6.61 | 6.0399 | 5.9512 | 5.98 | 6.024 | 6.72 | 63
| 1 | 6 | 5.78 | 5.91 | 6.53 | 6.88 | 6.02 | 6.1877 | 6.5 | 5.75 | 6.71 | 6.1696 | 5.9214 | 6.27 | 5.658 | 6.98
| Mean |  | 5.6083 | 5.95 | 6.285 | 6.7083 | 6.1467 | 6.3127 | 6.5033 | 5.7067 | 6.7 | 6.0993 | 5.955 | 6.0967 | 5.8595 | 6.855
| Median |  | 5.805 | 5.975 | 6.29 | 6.7 | 6.15 | 6.3241 | 6.51 | 5.555 | 6.715 | 6.1048 | 5.9608 | 6.1 | 5.9065 | 6.845
| Std Dev. |  | 0.45261 | 0.084617 | 0.19118 | 0.12703 | 0.096885 | 0.084753 | 0.035024 | 0.5507 | 0.056569 | 0.091695 | 0.043992 | 0.15655 | 0.19707 | 0.094181
| Rel.Std.Dev. |  | 8.07% | 1.422% | 3.042% | 1.894% | 1.576% | 1.343% | 0.5386% | 9.65% | 0.8443% | 1.503% | 0.7387% | 2.568% | 3.363% | 1.374%
| PDM3 |  | -9.676% | -4.174% | 1.222% | 8.039% | -1.006% | 1.667% | 4.738% | -8.093% | 7.905% | -1.77% | -4.092% | -1.812% | -5.631% | 10.4%
| Z-Score (Absolute) |  | 1.11 | 0.51 | 0.48 | 1.74 | 0.07 | 0.56 | 1.13 | 1.24 | 1.72 | 0.07 | 0.5 | 0.08 | 0.78 | 2.18
| Analytical results for As in OREAS 465 (Indicative Value 31.7 ppm) | [15 empty] | Indicative
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
1
| 1 | 1 | 25.3 | 24.6 | 14.2 | 23.5 | 47.4 | 25.8532 | 37 | 39 | 35.8 | 78.2 | < 50 | 74 |  |  | 1
| 1 | 2 | 24.2 | 27.3 | 12 | 22.9 | 46.8 | 26.1865 | 37 | 34 | 33.6 | 72.5 | < 50 | 74 |  |  | 16
| 1 | 3 | 22.9 | 25 | 11.9 | 24.6 | 43.8 | 25.7287 | 38 | 29 | 34.7 | 72 | < 50 | 76 |  |  | 16
| 1 | 4 | 23.6 | 24.7 | 13.6 | 21.9 | 45.4 | 25.5841 | 38 | 46 | 34.7 | 73.9 | < 50 | 77 |  |  | 31.7468
| 1 | 5 | 23.3 | 25.7 | 15 | 24.4 | 45.8 | 26.4 | 37 | 36 | 34.6 | 96.8 | < 50 | 77 |  |  | 29
| 1 | 6 | 26.7 | 27.9 | 14.4 | 26.5 | 43.9 | 25.8916 | 36 | 33 | 36.7 | 95.7 | < 50 | 75
| Mean |  | 24.3333 | 25.8667 | 13.5167 | 23.9667 | 45.5167 | 25.9407 | 37.1667 | 36.1667 | 35.0167 | 81.5167 |  | 75.5
| Median |  | 23.9 | 25.35 | 13.9 | 23.95 | 45.6 | 25.8724 | 37 | 35 | 34.7 | 76.05 |  | 75.5
| Std Dev. |  | 1.4292 | 1.4095 | 1.2937 | 1.5895 | 1.473 | 0.30118 | 0.75277 | 5.8452 | 1.0797 | 11.6244 |  | 1.3784
| Rel.Std.Dev. |  | 5.874% | 5.449% | 9.571% | 6.632% | 3.236% | 1.161% | 2.025% | 16.16% | 3.083% | 14.26% |  | 1.826%
| PDM3 |  | -23.35% | -18.52% | -57.42% | -24.51% | 43.37% | -18.29% | 17.07% | 13.92% | 10.3% | 156.8% |  | 137.8%
| Z-Score (Absolute) |  | 0.69 | 0.59 | 1.38 | 0.71 | 0.67 | 0.58 | 0.14 | 0.07 | 0 | 2.99 | NA | 2.6
| Analytical results for Ba in OREAS 465 (Certified Value 4359 ppm) | [15 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 0
0
| 1 | 1 | 4400 | 4420 | 4540 | 4740 | 4280 | 4387.733 | 4062 | 4377 | 4570 | 4216 | 4221.3 | 3917 | 4390 | 4170 | 1
| 1 | 2 | 4470 | 4280 | 4380 | 4570 | 4430 | 4422.5 | 4103 | 4382 | 4650 | 4113.6 | 4246.16 | 4064 | 4380 | 4160 | 2
| 1 | 3 | 4340 | 4410 | 4410 | 4700 | 4410 | 4369.6651 | 4081 | 4429 | 4570 | 4114 | 4284.95 | 3917 | 4375 | 4160 | 16
| 1 | 4 | 4470 | 4490 | 4590 | 4680 | 4450 | 4440.0524 | 4083 | 4417 | 4450 | 4223.8 | 4199.35 | 3830 | 4370 | 4310 | 4359.1998
| 1 | 5 | 4370 | 4440 | 4510 | 4310 | 4360 | 4464.8866 | 4080 | 4406 | 4420 | 4129.6 | 4262.44 | 3893 | 4380 | 4340 | 64
| 1 | 6 | 4340 | 4340 | 4630 | 4450 | 4290 | 4395.9193 | 4117 | 4394 | 4550 | 4133.2 | 4201.43 | 3784 | 4365 | 4150
| Mean |  | 4398.3333 | 4396.6667 | 4510 | 4575 | 4370 | 4413.4594 | 4087.6667 | 4400.8333 | 4535 | 4155.0333 | 4235.9383 | 3900.8333 | 4376.6667 | 4215
| Median |  | 4385 | 4415 | 4525 | 4625 | 4385 | 4409.2097 | 4082 | 4400 | 4560 | 4131.4 | 4233.73 | 3905 | 4377.5 | 4165
| Std Dev. |  | 59.8052 | 75.0111 | 98.5901 | 167.1825 | 72.3878 | 35.5559 | 19.3873 | 20.2526 | 85.2643 | 50.9309 | 34.5044 | 95.786 | 8.756 | 85.9651
| Rel.Std.Dev. |  | 1.36% | 1.706% | 2.186% | 3.654% | 1.656% | 0.8056% | 0.4743% | 0.4602% | 1.88% | 1.226% | 0.8146% | 2.456% | 0.2001% | 2.04%
| PDM3 |  | 0.8977% | 0.8595% | 3.459% | 4.95% | 0.2478% | 1.245% | -6.229% | 0.9551% | 4.033% | -4.684% | -2.828% | -10.51% | 0.4007% | -3.308%
| Z-Score (Absolute) |  | 0.06 | 0.05 | 0.61 | 0.93 | 0.08 | 0.13 | 1.48 | 0.07 | 0.74 | 1.15 | 0.75 | 2.41 | 0.05 | 0.85
| Analytical results for Be in OREAS 465 (Certified Value 11.6 ppm) | [15 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 10.5 | 10.8 | 10.9 | 12.95 | 7.88 | 11.54 | 10.9 | 12.2 | 11.3 | 12.08 | 10.9556 | 13.5 | 12 |  | 1
| 1 | 2 | 11 | 11.25 | 10.15 | 12.4 | 7.77 | 11.7249 | 11 | 13 | 11.6 | 11.31 | 11.1578 | 14 | 12.5 |  | 0
| 1 | 3 | 10.65 | 10.75 | 10.5 | 13 | 7.1 | 11.4823 | 10.9 | 13 | 11.4 | 11.8 | 11.1427 | 13.5 | 12 |  | 16
| 1 | 4 | 10.8 | 10.6 | 12.25 | 12.95 | 6.73 | 11.8106 | 11.1 | 13 | 11.8 | 12.23 | 11.0406 | 13.5 | 12.5 |  | 11.6226
| 1 | 5 | 11.15 | 11.1 | 12.7 | 12.2 | 8.36 | 11.6984 | 10.9 | 12.3 | 11.6 | 11.66 | 10.9013 | 14.5 | 13.2 |  | 65
| 1 | 6 | 12 | 10.65 | 11.15 | 12.55 | 8.2 | 11.9903 | 11.1 | 12 | 11.5 | 11.93 | 10.9897 | 13.5 | 11.9
| Mean |  | 11.0167 | 10.8583 | 11.275 | 12.675 | 7.6733 | 11.7077 | 10.9833 | 12.5833 | 11.5333 | 11.835 | 11.0313 | 13.75 | 12.35
| Median |  | 10.9 | 10.775 | 11.025 | 12.75 | 7.825 | 11.7116 | 10.95 | 12.65 | 11.55 | 11.865 | 11.0151 | 13.5 | 12.25
| Std Dev. |  | 0.53541 | 0.25965 | 1.0004 | 0.33875 | 0.63554 | 0.18425 | 0.098319 | 0.46655 | 0.17512 | 0.32636 | 0.10281 | 0.41833 | 0.49295
| Rel.Std.Dev. |  | 4.86% | 2.391% | 8.873% | 2.673% | 8.282% | 1.574% | 0.8952% | 3.708% | 1.518% | 2.758% | 0.932% | 3.042% | 3.992%
| PDM3 |  | -5.214% | -6.576% | -2.991% | 9.054% | -33.98% | 0.7322% | -5.501% | 8.266% | -0.7684% | 1.827% | -5.088% | 18.3% | 6.258%
| Z-Score (Absolute) |  | 0.63 | 0.83 | 0.32 | 1.4 | 4.73 | 0.21 | 0.67 | 1.29 | 0 | 0.37 | 0.62 | 2.53 | 1
| Analytical results for Bi in OREAS 465 (Certified Value 17.3 ppm) | [15 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 17.7 | 16.3 | 17.3 | 17.3 | 17.05 | 16.7402 | 17.7 | 19.2 | < 0.01 | 17.82 | 18.0036 | 16.6 | 16.4 |  | 1
| 1 | 2 | 17.9 | 16.3 | 17.1 | 17.3 | 17.25 | 15.9835 | 17.9 | 18.7 | < 0.01 | 17.71 | 18.0662 | 16.4 | 16.2 |  | 0
| 1 | 3 | 17.65 | 16.4 | 17.2 | 18 | 16.65 | 16.0457 | 17.8 | 19 | < 0.01 | 18.38 | 18.0665 | 16.3 | 16.3 |  | 16
| 1 | 4 | 18.05 | 16.55 | 18.95 | 17.9 | 16.65 | 15.9745 | 17.7 | 19.6 | < 0.01 | 18.71 | 17.4265 | 15.8 | 16.8 |  | 17.2678
| 1 | 5 | 17.7 | 16.55 | 18.9 | 15.15 | 15.65 | 16.3043 | 18.1 | 19 | < 0.01 | 18.05 | 17.7023 | 16.2 | 17.2 |  | 66
| 1 | 6 | 16.35 | 16.5 | 18.35 | 15.4 | 15.35 | 16.929 | 17.9 | 18.8 | < 0.01 | 17.36 | 17.6403 | 16.3 | 16.4
| Mean |  | 17.5583 | 16.4333 | 17.9667 | 16.8417 | 16.4333 | 16.3295 | 17.85 | 19.05 |  | 18.005 | 17.8176 | 16.2667 | 16.55
| Median |  | 17.7 | 16.45 | 17.825 | 17.3 | 16.65 | 16.175 | 17.85 | 19 |  | 17.935 | 17.853 | 16.3 | 16.4
| Std Dev. |  | 0.61108 | 0.1169 | 0.86814 | 1.2508 | 0.76529 | 0.41351 | 0.15166 | 0.32094 |  | 0.48517 | 0.26686 | 0.26583 | 0.37815
| Rel.Std.Dev. |  | 3.48% | 0.7114% | 4.832% | 7.427% | 4.657% | 2.532% | 0.8496% | 1.685% |  | 2.695% | 1.498% | 1.634% | 2.285%
| PDM3 |  | 1.682% | -4.833% | 4.047% | -2.468% | -4.833% | -5.434% | 3.372% | 10.32% |  | 4.269% | 3.184% | -5.798% | -4.157%
| Z-Score (Absolute) |  | 0.41 | 0.76 | 0.55 | 0.41 | 0.76 | 0.85 | 0.45 | 1.48 | NA | 0.59 | 0.43 | 0.9 | 0.77
| Analytical results for Ca in OREAS 465 (Certified Value 0.872 wt.%) | [15 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.86 | 0.79 | 0.42 | 0.51 | 0.41 | 0.88502 | 0.84 | 0.95 | 0.774 | 0.9198 | 0.82079 | 0.88 | 0.896 | 0.97 | 1
| 1 | 2 | 0.86 | 0.8 | 0.42 | 0.48 | 0.42 | 0.9251 | 0.85 | 0.93 | 0.768 | 0.9358 | 0.82134 | 0.87 | 0.896 | 0.95 | 3
| 1 | 3 | 0.85 | 0.79 | 0.41 | 0.5 | 0.43 | 0.92425 | 0.84 | 0.95 | 0.767 | 0.9198 | 0.82045 | 0.87 | 0.897 | 0.98 | 16
| 1 | 4 | 0.85 | 0.8 | 0.41 | 0.49 | 0.42 | 0.92636 | 0.84 | 0.95 | 0.777 | 0.9262 | 0.81923 | 0.87 | 0.901 | 0.97 | 0.87181
| 1 | 5 | 0.84 | 0.81 | 0.41 | 0.78 | 0.42 | 0.84947 | 0.85 | 0.97 | 0.768 | 0.9032 | 0.82962 | 0.88 | 0.8945 | 0.96 | 67
| 1 | 6 | 0.86 | 0.79 | 0.41 | 0.81 | 0.41 | 0.82929 | 0.84 | 0.95 | 0.773 | 0.9237 | 0.82703 | 0.85 | 0.9115 | 1
| Mean |  | 0.85333 | 0.79667 | 0.41333 | 0.595 | 0.41833 | 0.88992 | 0.84333 | 0.95 | 0.77117 | 0.92142 | 0.82308 | 0.87 | 0.89933 | 0.97167
| Median |  | 0.855 | 0.795 | 0.41 | 0.505 | 0.42 | 0.90464 | 0.84 | 0.95 | 0.7705 | 0.92175 | 0.82106 | 0.87 | 0.8965 | 0.97
| Std Dev. |  | 0.008165 | 0.008165 | 0.005164 | 0.15553 | 0.0075277 | 0.042614 | 0.005164 | 0.012649 | 0.0040702 | 0.010696 | 0.0042074 | 0.010954 | 0.0063535 | 0.017224
| Rel.Std.Dev. |  | 0.9568% | 1.025% | 1.249% | 26.14% | 1.799% | 4.789% | 0.6123% | 1.331% | 0.5278% | 1.161% | 0.5112% | 1.259% | 0.7065% | 1.773%
| PDM3 |  | -2.119% | -8.619% | -52.59% | -31.75% | -52.02% | 2.077% | -3.266% | 8.969% | -11.54% | 5.69% | -5.59% | -0.2076% | 3.157% | 11.45%
| Z-Score (Absolute) |  | 0.05 | 0.56 | 4.7 | 2.74 | 4.65 | 0.45 | 0.05 | 1.1 | 0.83 | 0.79 | 0.27 | 0.23 | 0.55 | 1.33
| Analytical results for Cd in OREAS 465 (Certified Value 1.2 ppm) | [15 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 2
2
| 1 | 1 | 1.18 | 1.14 | 1.16 | 1.19 | 1.24 | 1.2901 | 1 | 1.75 | 0.95 | 1 | 1 |  |  |  | 1
| 1 | 2 | 1.24 | 1.23 | 1.12 | 1.29 | 1.28 | 1.3255 | 1.2 | 1.87 | 0.95 | 1 | 1.5 |  |  |  | 18
| 1 | 3 | 1.18 | 1.11 | 1.24 | 1.26 | 1.21 | 1.2702 | 1.4 | 1.89 | 0.91 | 1 | 1 |  |  |  | 16
| 1 | 4 | 1.2 | 1.09 | 1.3 | 1.28 | 1.24 | 1.421 | 1.2 | 1.87 | 0.86 | 1 | 1.5 |  |  |  | 1.203
| 1 | 5 | 1.16 | 1.17 | 1.31 | 1.2 | 1.12 | 1.2658 | 1 | 1.81 | 0.83 | 1 | 1.5 |  |  |  | 68
| 1 | 6 | 1.25 | 1.23 | 1.2 | 1.24 | 1.12 | 1.4121 | 1 | 1.84 | 0.92 | 1 | 1.5
| Mean |  | 1.2017 | 1.1617 | 1.2217 | 1.2433 | 1.2017 | 1.3308 | 1.1333 | 1.8383 | 0.90333 | 1 | 1.3333
| Median |  | 1.19 | 1.155 | 1.22 | 1.25 | 1.225 | 1.3078 | 1.1 | 1.855 | 0.915 | 1 | 1.5
| Std Dev. |  | 0.036009 | 0.05947 | 0.076004 | 0.041312 | 0.067057 | 0.06976 | 0.1633 | 0.051543 | 0.048854 | 0 | 0.2582
| Rel.Std.Dev. |  | 2.997% | 5.119% | 6.221% | 3.323% | 5.58% | 5.242% | 14.41% | 2.804% | 5.408% | 0% | 19.36%
| PDM3 |  | -0.1149% | -3.44% | 1.548% | 3.349% | -0.1149% | 10.62% | -5.795% | 52.81% | -24.91% | -16.88% | 10.83%
| Z-Score (Absolute) |  | 0 | 0.39 | 0.2 | 0.41 | 0 | 1.27 | 0.67 | 6.28 | 2.94 | 1.99 | 1.3
| Analytical results for Ce in OREAS 465 (Certified Value 3.91 wt.%) | [15 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 4.0492 | 3.7765 | 3.6721 | > 0.5 | 3.8002 | 4.0906 | 3.7813 | 3.89 |  | 1
| 1 | 2 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 4.0469 | 3.8316 | 3.6379 | > 0.5 | 3.8698 | 4.0654 | 3.8554 | 3.89 |  | 0
| 1 | 3 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 4.0557 | 3.7669 | 3.6442 | > 0.5 | 3.9336 | 4.2399 | 3.9028 | 3.91 |  | 16
| 1 | 4 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 4.0062 | 3.7712 | 3.8714 | > 0.5 | 4.0064 | 4.1186 | 3.8607 | 4.02 |  | 3.9056
| 1 | 5 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 4.0591 | 3.8207 | 3.704 | > 0.5 | 3.8899 | 4.1074 | 3.9539 | 4.07 |  | 69
| 1 | 6 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | > 0.05 | 4.0168 | 3.8552 | 3.6426 | > 0.5 | 3.8222 | 4.1359 | 3.9051 | 3.9
| Mean |  | 07: 4.039; 08: 3.8037; 09: 3.6954; 11: 3.887; 12: 4.1263; 17: 3.8765; 22: 3.9467
| Median |  | 07: 4.0481; 08: 3.7986; 09: 3.6582; 11: 3.8799; 12: 4.113; 17: 3.8817; 22: 3.905
| Std Dev. |  | 07: 0.021992; 08: 0.037073; 09: 0.089768; 11: 0.075421; 12: 0.060659; 17: 0.058761; 22: 0.078145
| Rel.Std.Dev. |  | 07: 0.5445%; 08: 0.9747%; 09: 2.429%; 11: 1.94%; 12: 1.47%; 17: 1.516%; 22: 1.98%
| PDM3 |  | 07: 3.415%; 08: -2.61%; 09: -5.384%; 11: -0.4762%; 12: 5.65%; 17: -0.7444%; 22: 1.051%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 1.23 | 0.67 | 1.84 | NA | 0 | 1.94 | 0.08 | 0.48
| Analytical results for Co in OREAS 465 (Certified Value 18.7 ppm) | [15 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 18.6 | 18.3 | 18.7 | 19.4 | 18 | 18.4328 | 19.6 | 18 | 18.4 | 17.3 | 19.9961 | 45 | 15 |  | 1
| 1 | 2 | 19 | 18.8 | 18.4 | 19.4 | 18.5 | 18.1763 | 20.1 | 18 | 18.5 | 17.5 | 20.9866 | 45 | 15 |  | 19
| 1 | 3 | 17.4 | 18.1 | 18 | 20.3 | 17.3 | 18.2206 | 19.8 | 19 | 18.8 | 17.7 | 20.3954 | 35 | 20 |  | 16
| 1 | 4 | 18.8 | 18.5 | 19.7 | 20.1 | 17.6 | 18.1827 | 19.9 | 20 | 18.8 | 17.5 | 21.3254 | 30 | 25 |  | 18.7017
| 1 | 5 | 17.9 | 18.5 | 19.7 | 18.5 | 18.1 | 18.1155 | 19.6 | 22 | 19.1 | 18.4 | 19.6664 | 30 | 20 |  | 70
| 1 | 6 | 19.7 | 18.2 | 18 | 18.9 | 18.1 | 18.3937 | 20.2 | 19 | 18.4 | 18.4 | 20.13 | 35 | 15
| Mean |  | 18.5667 | 18.4 | 18.75 | 19.4333 | 17.9333 | 18.2536 | 19.8667 | 19.3333 | 18.6667 | 17.8 | 20.4167 | 36.6667 | 18.3333
| Median |  | 18.7 | 18.4 | 18.55 | 19.4 | 18.05 | 18.2017 | 19.85 | 19 | 18.65 | 17.6 | 20.2627 | 35 | 17.5
| Std Dev. |  | 0.8165 | 0.25298 | 0.78166 | 0.68605 | 0.42269 | 0.12874 | 0.25033 | 1.5055 | 0.28048 | 0.48166 | 0.62816 | 6.8313 | 4.0825
| Rel.Std.Dev. |  | 4.398% | 1.375% | 4.169% | 3.53% | 2.357% | 0.7053% | 1.26% | 7.787% | 1.503% | 2.706% | 3.077% | 18.63% | 22.27%
| PDM3 |  | -0.722% | -1.613% | 0.2583% | 3.912% | -4.108% | -2.396% | 6.229% | 3.377% | -0.1873% | -4.821% | 9.17% | 96.06% | -1.97%
| Z-Score (Absolute) |  | 0.09 | 0.25 | 0.08 | 0.7 | 0.67 | 0.38 | 1.1 | 0.61 | 0 | 0.8 | 1.61 | 16.55 | 1.53
| Analytical results for Cr in OREAS 465 (Indicative Value 357 ppm) | [15 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  |  | 0
0
| 1 | 1 | 343 | 325 | 189 | 210 | 179 | 449.5105 | 404 | 482 | 420 | 306 | 600 |  |  |  | 1
| 1 | 2 | 347 | 317 | 179 | 203 | 185 | 459.6601 | 382 | 479 | 426 | 292 | 550 |  |  |  | 4
| 1 | 3 | 319 | 322 | 175 | 212 | 179 | 456.2652 | 423 | 489 | 414 | 283 | 600 |  |  |  | 16
| 1 | 4 | 348 | 333 | 196 | 206 | 186 | 452.6634 | 435 | 476 | 434 | 290 | 600 |  |  |  | 356.9204
| 1 | 5 | 337 | 322 | 189 | 362 | 188 | 461.7849 | 448 | 492 | 412 | 277 | 550 |  |  |  | 30
| 1 | 6 | 326 | 325 | 190 | 373 | 172 | 459.6622 | 388 | 467 | 484 | 290 | 550
| Mean |  | 336.6667 | 324 | 186.3333 | 261 | 181.5 | 456.591 | 413.3333 | 480.8333 | 431.6667 | 289.6667 | 575
| Median |  | 340 | 323.5 | 189 | 211 | 182 | 457.9627 | 413.5 | 480.5 | 423 | 290 | 575
| Std Dev. |  | 11.8434 | 5.2915 | 7.7889 | 82.6269 | 5.9582 | 4.7222 | 26.3641 | 9.0646 | 26.875 | 9.7707 | 27.3861
| Rel.Std.Dev. |  | 3.518% | 1.633% | 4.18% | 31.66% | 3.283% | 1.034% | 6.378% | 1.885% | 6.226% | 3.373% | 4.763%
| PDM3 |  | -5.675% | -9.223% | -47.79% | -26.87% | -49.15% | 27.93% | 15.81% | 34.72% | 20.94% | -18.84% | 61.1%
| Z-Score (Absolute) |  | 0 | 0.1 | 1.2 | 0.6 | 1.24 | 0.96 | 0.61 | 1.15 | 0.67 | 0.37 | 1.9
| Analytical results for Cs in OREAS 465 (Certified Value < 0.1 ppm) | [15 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.5 | 0.05 | < 0.05 | 0.18045 | < 0.5 | < 0.1 |  | 1
| 1 | 2 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.5 | < 0.05 | < 0.05 | 0.17221 | < 0.5 | < 0.1 |  | 0
| 1 | 3 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.5 | < 0.05 | < 0.05 | 0.15882 | < 0.5 | < 0.1 |  | 16
| 1 | 4 | < 0.05 | < 0.05 | 0.05 | < 0.05 | < 0.05 | < 0.05 | < 0.1 | < 0.5 | 0.05 | < 0.05 | 0.14693 | < 0.5 | < 0.1 |  | < 0.1
| 1 | 5 | < 0.05 | < 0.05 | 0.06 | 0.06 | < 0.05 | < 0.05 | < 0.1 | < 0.5 | 0.05 | < 0.05 | 0.15681 | < 0.5 | < 0.1 |  | 71
| 1 | 6 | < 0.05 | < 0.05 | < 0.05 | 0.06 | < 0.05 | < 0.05 | < 0.1 | < 0.5 | 0.05 | < 0.05 | 0.14057 | < 0.5 | < 0.1
| Mean |  | 03: 0.055; 05: 0.06; 10: 0.05; 12: 0.1593
| Median |  | 03: 0.055; 05: 0.06; 10: 0.05; 12: 0.15781
| Std Dev. |  | 03: 0.0070711; 05: 0; 10: 0; 12: 0.014996
| Rel.Std.Dev. |  | 03: 12.86%; 05: 0%; 10: 0%; 12: 9.414%
PDM3
| Z-Score (Absolute) |  | NA | NA | 1.06 | 0.29 | NA | NA | NA | NA | 0.29 | NA | 13.89 | NA | NA
| Analytical results for Cu in OREAS 465 (Certified Value 128 ppm) | [15 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 129 | 124 | 124 | 125.5 | 124 | 135.6949 | 136.2 | 149 | 124 | 140 | 128.58 | 104 | 130 | 120 | 1
| 1 | 2 | 130.5 | 121 | 118.5 | 125 | 126 | 134.8436 | 138.7 | 132 | 127 | 113 | 128.404 | 106 | 130 | 126 | 20
| 1 | 3 | 121.5 | 122 | 120.5 | 130.5 | 120.5 | 133.882 | 144.3 | 147 | 126 | 121 | 127.856 | 100 | 130 | 120 | 16
| 1 | 4 | 129.5 | 125 | 134.5 | 130 | 122 | 132.2442 | 134.6 | 151 | 128 | 141 | 128.529 | 103 | 140 | 130 | 127.5667
| 1 | 5 | 124 | 126 | 132.5 | 123.5 | 124.5 | 132.1994 | 130.1 | 157 | 129 | 131 | 128.166 | 99.5 | 130 | 128 | 72
| 1 | 6 | 125 | 124 | 128 | 132 | 122 | 133.7796 | 139.8 | 134 | 127 | 126 | 127.223 | 105 | 130 | 136
| Mean |  | 126.5833 | 123.6667 | 126.3333 | 127.75 | 123.1667 | 133.7739 | 137.2833 | 145 | 126.8333 | 128.6667 | 128.1263 | 102.9167 | 131.6667 | 126.6667
| Median |  | 127 | 124 | 126 | 127.75 | 123 | 133.8308 | 137.45 | 148 | 127 | 128.5 | 128.285 | 103.5 | 130 | 127
| Std Dev. |  | 3.5975 | 1.8619 | 6.455 | 3.5036 | 2.0166 | 1.3904 | 4.8487 | 9.8995 | 1.7224 | 10.9301 | 0.51686 | 2.6536 | 4.0825 | 6.1536
| Rel.Std.Dev. |  | 2.842% | 1.506% | 5.109% | 2.743% | 1.637% | 1.039% | 3.532% | 6.827% | 1.358% | 8.495% | 0.4034% | 2.578% | 3.101% | 4.858%
| PDM3 |  | -0.7709% | -3.057% | -0.9668% | 0.1437% | -3.449% | 4.866% | 7.617% | 13.67% | -0.5749% | 0.8623% | 0.4387% | -19.32% | 3.214% | -0.7055%
| Z-Score (Absolute) |  | 0.19 | 0.98 | 0.26 | 0.12 | 1.11 | 1.75 | 2.69 | 4.78 | 0.12 | 0.37 | 0.23 | 6.57 | 1.18 | 0.17
| Analytical results for Dy in OREAS 465 (Certified Value 215 ppm) | [15 empty] | Certified
Dy | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 210 | 202 | 207 | 211 | 216 | 219.1764 | 212.9 | 219.2 | 162 | 225.69 | 219.2454 | 209.5 | 231 |  | 1
| 1 | 2 | 210 | 193 | 196.5 | 209 | 217 | 223.1162 | 227.4 | 217 | 167 | 218.07 | 218.8535 | 209.4 | 229 |  | 5
| 1 | 3 | 196 | 207 | 203 | 219 | 208 | 219.0496 | 214.9 | 219 | 162 | 226.42 | 220.054 | 208.9 | 225 |  | 16
| 1 | 4 | 213 | 199.5 | 210 | 218 | 209 | 212.7943 | 218.2 | 228 | 169 | 227.74 | 220.0069 | 210.2 | 228 |  | 215.1915
| 1 | 5 | 196.5 | 206 | 212 | 218 | 224 | 216.081 | 234.3 | 217.2 | 161 | 222.8 | 218.8419 | 213.4 | 233 |  | 73
| 1 | 6 | 197 | 209 | 201 | 220 | 216 | 223.3234 | 220.4 | 215.3 | 168 | 223.03 | 216.2537 | 211 | 228
| Mean |  | 203.75 | 202.75 | 204.9167 | 215.8333 | 215 | 218.9235 | 221.35 | 219.2833 | 164.8333 | 223.9583 | 218.8759 | 210.4 | 229
| Median |  | 203.5 | 204 | 205 | 218 | 216 | 219.113 | 219.3 | 218.1 | 164.5 | 224.36 | 219.0494 | 209.85 | 228.5
| Std Dev. |  | 8.0234 | 5.897 | 5.8345 | 4.6224 | 5.8652 | 4.0661 | 8.101 | 4.5053 | 3.5449 | 3.4709 | 1.3915 | 1.6407 | 2.7568
| Rel.Std.Dev. |  | 3.938% | 2.909% | 2.847% | 2.142% | 2.728% | 1.857% | 3.66% | 2.055% | 2.151% | 1.55% | 0.6357% | 0.7798% | 1.204%
| PDM3 |  | -5.317% | -5.782% | -4.775% | 0.2983% | -0.08898% | 1.734% | 2.862% | 1.901% | -23.4% | 4.074% | 1.712% | -2.227% | 6.417%
| Z-Score (Absolute) |  | 1.48 | 1.6 | 1.33 | 0 | 0.1 | 0.38 | 0.67 | 0.21 | 6.23 | 0.99 | 0.37 | 0.66 | 1.61
| Analytical results for Er in OREAS 465 (Certified Value 47.3 ppm) | [15 empty] | Certified
Er | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 49 | 43.9 | 52.6 | 44.7 | 48.4 | 48.1875 | 45.3 | 46.9 | 49.8 | 45.41 | 79.2803 | 46.1 | 45.3 |  | 1
| 1 | 2 | 49.3 | 41.2 | 49.1 | 44.9 | 49.1 | 48.5765 | 45.9 | 46.6 | 51 | 45.4 | 77.8211 | 45.5 | 43.7 |  | 6
| 1 | 3 | 45.5 | 45.4 | 50.8 | 46.8 | 47.2 | 50.3963 | 44.6 | 47 | 50 | 47.4 | 81.0453 | 45.65 | 44.3 |  | 16
| 1 | 4 | 49.6 | 44.8 | 54.1 | 46.6 | 47.1 | 48.9 | 44.3 | 47.4 | 48.2 | 48.23 | 75.9968 | 46.4 | 45.1 |  | 47.265
| 1 | 5 | 45.6 | 45.4 | 54.4 | 45.2 | 50.9 | 49.3935 | 45.8 | 46 | 50.1 | 48.38 | 76.6549 | 47 | 46.1 |  | 74
| 1 | 6 | 43.5 | 45.1 | 52 | 45.7 | 50 | 48.8969 | 45.2 | 45.6 | 50 | 45.16 | 71.355 | 46.3 | 44
| Mean |  | 47.0833 | 44.3 | 52.1667 | 45.65 | 48.7833 | 49.0585 | 45.1833 | 46.5833 | 49.85 | 46.6633 | 77.0256 | 46.1583 | 44.75
| Median |  | 47.3 | 44.95 | 52.3 | 45.45 | 48.75 | 48.8985 | 45.25 | 46.75 | 50 | 46.405 | 77.238 | 46.2 | 44.7
| Std Dev. |  | 2.5483 | 1.6174 | 2.0106 | 0.88261 | 1.5198 | 0.76737 | 0.63692 | 0.67057 | 0.91159 | 1.508 | 3.3234 | 0.54444 | 0.90719
| Rel.Std.Dev. |  | 5.412% | 3.651% | 3.854% | 1.933% | 3.115% | 1.564% | 1.41% | 1.44% | 1.829% | 3.232% | 4.315% | 1.18% | 2.027%
| PDM3 |  | -0.3844% | -6.273% | 10.37% | -3.417% | 3.212% | 3.794% | -4.404% | -1.442% | 5.469% | -1.273% | 62.97% | -2.341% | -5.321%
| Z-Score (Absolute) |  | 0.16 | 0.67 | 2.13 | 0.39 | 0.82 | 0.93 | 0.57 | 0.03 | 1.36 | 0 | 11.74 | 0.2 | 0.74
| Analytical results for Eu in OREAS 465 (Certified Value 282 ppm) | [15 empty] | Certified
Eu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 324 | 274 | 288 | 300 | 292 | 282.4597 | 290.2 | 282.4 | 129 | 283.3 | 284.3081 | 253.7 | 262 |  | 1
| 1 | 2 | 331 | 264 | 273 | 298 | 295 | 285.5482 | 293.6 | 279.2 | 130 | 269.88 | 284.4017 | 262.1 | 262 |  | 7
| 1 | 3 | 302 | 271 | 275 | 310 | 289 | 287.9432 | 298.6 | 278.2 | 128 | 288.41 | 285.4658 | 263.3 | 265 |  | 16
| 1 | 4 | 329 | 280 | 290 | 308 | 290 | 283.8 | 286 | 298.1 | 127 | 289.35 | 285.1785 | 262.5 | 271 |  | 282.2001
| 1 | 5 | 307 | 274 | 289 | 295 | 299 | 283.7793 | 288.1 | 289 | 124 | 277.14 | 284.3539 | 267.1 | 272 |  | 75
| 1 | 6 | 299 | 273 | 280 | 295 | 291 | 289.7745 | 296.6 | 282 | 126 | 278.81 | 286.6465 | 265.5 | 259
| Mean |  | 315.3333 | 272.6667 | 282.5 | 301 | 292.6667 | 285.5508 | 292.1833 | 284.8167 | 127.3333 | 281.1483 | 285.0591 | 262.3667 | 265.1667
| Median |  | 315.5 | 273.5 | 284 | 299 | 291.5 | 284.6741 | 291.9 | 282.2 | 127.5 | 281.055 | 284.7901 | 262.9 | 263.5
| Std Dev. |  | 14.2922 | 5.2026 | 7.5033 | 6.5115 | 3.7238 | 2.8044 | 4.9301 | 7.5239 | 2.1602 | 7.3911 | 0.91574 | 4.6539 | 5.2694
| Rel.Std.Dev. |  | 4.532% | 1.908% | 2.656% | 2.163% | 1.272% | 0.9821% | 1.687% | 2.642% | 1.697% | 2.629% | 0.3212% | 1.774% | 1.987%
| PDM3 |  | 11.74% | -3.378% | 0.1063% | 6.662% | 3.709% | 1.187% | 3.538% | 0.9272% | -54.88% | -0.3727% | 1.013% | -7.028% | -6.036%
| Z-Score (Absolute) |  | 2.25 | 0.67 | 0 | 1.27 | 0.7 | 0.21 | 0.66 | 0.02 | 10.64 | 0.09 | 0.18 | 1.26 | 1.19
| Analytical results for Fe in OREAS 465 (Certified Value 29.55 wt.%) | [15 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 27 | 27.8 | 28 | 30.6 | 26.8 | 31.645 | 30.19 | 26.64 | 32.3 | 30.7 | 28.5646 | > 15 | 30.2925 | 32.3 | 1
| 1 | 2 | 27.5 | 26.9 | 27.3 | 29.5 | 27.3 | 31.6277 | 30.63 | 25.85 | 32.7 | 32.33 | 28.2006 | > 15 | 30.644 | 31.6 | 0
| 1 | 3 | 24.5 | 27.8 | 27.4 | 30.4 | 27.5 | 31.8553 | 30.38 | 27.69 | 32.6 | 31.66 | 28.5815 | - | 30.8655 | 31.9 | 16
| 1 | 4 | 27.6 | 28 | 28.7 | 29.9 | 27.6 | 31.5541 | 30.14 | 29.87 | 32.5 | 30.6 | 28.5434 | > 15 | 30.736 | 31.7 | 29.5527
| 1 | 5 | 25.8 | 28.1 | 28.2 | 28.7 | 27.3 | 31.7682 | 30.33 | 31.53 | 32.9 | 30.09 | 29.0375 | > 15 | 30.925 | 31.1 | 76
| 1 | 6 | 26.7 | 27.7 | 28.9 | 30 | 26.6 | 31.568 | 30.23 | 26.94 | 32 | 30.51 | 28.4759 | > 15 | 30.733 | 32.5
| Mean |  | 26.5167 | 27.7167 | 28.0833 | 29.85 | 27.1833 | 31.6697 | 30.3167 | 28.0867 | 32.5 | 30.9817 | 28.5673 |  | 30.6993 | 31.85
| Median |  | 26.85 | 27.8 | 28.1 | 29.95 | 27.3 | 31.6364 | 30.28 | 27.315 | 32.55 | 30.65 | 28.554 |  | 30.7345 | 31.8
| Std Dev. |  | 1.1822 | 0.42622 | 0.65549 | 0.68337 | 0.39707 | 0.11849 | 0.17728 | 2.1728 | 0.31623 | 0.83913 | 0.27007 |  | 0.22342 | 0.50498
| Rel.Std.Dev. |  | 4.458% | 1.538% | 2.334% | 2.289% | 1.461% | 0.3742% | 0.5847% | 7.736% | 0.973% | 2.708% | 0.9454% |  | 0.7278% | 1.585%
| PDM3 |  | -10.27% | -6.213% | -4.972% | 1.006% | -8.017% | 7.164% | 2.585% | -4.961% | 9.973% | 4.835% | -3.334% |  | 3.88% | 7.774%
| Z-Score (Absolute) |  | 1.27 | 0.75 | 0.67 | 0 | 1.02 | 0.69 | 0.18 | 0.67 | 1.01 | 0.43 | 0.49 | NA | 0.32 | 0.76
| Analytical results for Ga in OREAS 465 (Certified Value 188 ppm) | [15 empty] | Certified
Ga | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 02; 02: 03; 03: 05; 05: 06; 06: 07; 07: 10; 08: 12; 09: 22 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS | 0
0
| 1 | 1 | 01: 221; 02: 157.5; 03: 180; 05: 283; 06: 88; 07: < 0.05; 08: 193.9365; 09: 166 | 1
| 1 | 2 | 01: 242; 02: 153.5; 03: 178.5; 05: 293; 06: 95.2589; 07: < 0.05; 08: 193.3214; 09: 177 | 0
| 1 | 3 | 01: 203; 02: 153; 03: 187; 05: 272; 06: 89.8292; 07: < 0.05; 08: 193.4279; 09: 149 | 16
| 1 | 4 | 01: 212; 02: 190.5; 03: 187; 05: 276; 06: 87.655; 07: < 0.05; 08: 191.086; 09: 143 | 188.3139
| 1 | 5 | 01: 212; 02: 186.5; 03: 266; 05: 286; 06: 91; 07: < 0.05; 08: 193.3072; 09: 155 | 77
| 1 | 6 | 01: 213; 02: 177; 03: 268; 05: 276; 06: 94; 07: < 0.05; 08: 192.6392; 09: 144
| Mean |  | 01: 217.1667; 02: 169.6667; 03: 211.0833; 05: 281; 06: 90.9572; 08: 192.953; 09: 155.6667
| Median |  | 01: 212.5; 02: 167.25; 03: 187; 05: 279.5; 06: 90.4146; 08: 193.3143; 09: 152
| Std Dev. |  | 01: 13.4375; 02: 17.0783; 03: 43.4585; 05: 7.7974; 06: 3.1205; 08: 1.0038; 09: 13.4412
| Rel.Std.Dev. |  | 01: 6.188%; 02: 10.07%; 03: 20.59%; 05: 2.775%; 06: 3.431%; 08: 0.5202%; 09: 8.635%
| PDM3 |  | 01: 15.32%; 02: -9.902%; 03: 12.09%; 05: 49.22%; 06: -51.7%; 08: 2.463%; 09: -17.34%
| Z-Score (Absolute) |  | 01: 0.56; 02: 0.67; 03: 0.53; 05: 2.55; 06: 2.95; 07: NA; 08: 0; 09: 1.08
| Analytical results for Gd in OREAS 465 (Certified Value 581 ppm) | [15 empty] | Certified
Gd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 601 | 557 | 826 | 521 | 652 | 579.2243 | 630.3 | 550.1 | 149 | 587.31 | 893.9606 | 536 | 573 |  | 1
| 1 | 2 | 592 | 558 | 758 | 523 | 664 | 575.5944 | 630.2 | 549.2 | 154 | 582.18 | 895.4532 | 548.4 | 567 |  | 8
| 1 | 3 | 545 | 566 | 799 | 545 | 636 | 581.92 | 619 | 552.3 | 155 | 613.15 | 893.7396 | 554.6 | 565 |  | 16
| 1 | 4 | 599 | 565 | 790 | 541 | 634 | 582.5299 | 614.8 | 587.9 | 157 | 605.68 | 892.5786 | 553.1 | 572 |  | 581.0732
| 1 | 5 | 551 | 570 | 787 | 545 | 678 | 585.5935 | 612.8 | 560.3 | 155 | 606.55 | 896.9571 | 559.6 | 581 |  | 78
| 1 | 6 | 559 | 575 | 784 | 545 | 658 | 581.9294 | 633.7 | 553.9 | 150 | 595.89 | 895.0137 | 557.7 | 561
| Mean |  | 574.5 | 565.1667 | 790.6667 | 536.6667 | 653.6667 | 581.1319 | 623.4667 | 558.95 | 153.3333 | 598.46 | 894.6171 | 551.5667 | 569.8333
| Median |  | 575.5 | 565.5 | 788.5 | 543 | 655 | 581.9247 | 624.6 | 553.1 | 154.5 | 600.785 | 894.4871 | 553.85 | 569.5
| Std Dev. |  | 25.5793 | 6.9113 | 22.1058 | 11.4833 | 16.8483 | 3.3888 | 9.0066 | 14.7176 | 3.1411 | 12.0776 | 1.529 | 8.5563 | 7.0545
| Rel.Std.Dev. |  | 4.452% | 1.223% | 2.796% | 2.14% | 2.578% | 0.5831% | 1.445% | 2.633% | 2.049% | 2.018% | 0.1709% | 1.551% | 1.238%
| PDM3 |  | -1.131% | -2.737% | 36.07% | -7.642% | 12.49% | 0.0101% | 7.296% | -3.807% | -73.61% | 2.992% | 53.96% | -5.078% | -1.934%
| Z-Score (Absolute) |  | 0 | 0.26 | 6.08 | 1.06 | 2.23 | 0.19 | 1.38 | 0.6 | 11.85 | 0.67 | 9.01 | 0.56 | 0.13
| Analytical results for Ge in OREAS 465 (Indicative Value 23.6 ppm) | [15 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 12.1 | 19.45 | 17.7 | 17.2 | 34.5 | 0.2 | 24.1 | 34.9 | < 0.05 | 25.34 |  |  |  |  | 1
| 1 | 2 | 12 | 22.8 | 16.85 | 17.3 | 35 | 0.2 | 24.5 | 34.4 | < 0.05 | 23.42 |  |  |  |  | 21
| 1 | 3 | 11.55 | 19.3 | 16.9 | 18.1 | 33.7 | 0.26059 | 24 | 33.5 | < 0.05 | 24.35 |  |  |  |  | 16
| 1 | 4 | 11.5 | 19.5 | 21.9 | 17.9 | 33.9 | 0.2 | 24.4 | 39.5 | < 0.05 | 25.6 |  |  |  |  | 23.5885
| 1 | 5 | 11.3 | 19.8 | 21.6 | 24.5 | 33.6 | 0.22911 | 23.4 | 36.8 | < 0.05 | 25.07 |  |  |  |  | 31
| 1 | 6 | 20.3 | 19.25 | 20.6 | 24 | 32.7 | 0.2 | 24.4 | 36.1 | < 0.05 | 23.62
| Mean |  | 13.125 | 20.0167 | 19.2583 | 19.8333 | 33.9 | 0.21495 | 24.1333 | 35.8667 |  | 24.5667
| Median |  | 11.775 | 19.475 | 19.15 | 18 | 33.8 | 0.2 | 24.25 | 35.5 |  | 24.71
| Std Dev. |  | 3.5284 | 1.3772 | 2.3686 | 3.4419 | 0.79246 | 0.025208 | 0.40825 | 2.136 |  | 0.91402
| Rel.Std.Dev. |  | 26.88% | 6.88% | 12.3% | 17.35% | 2.338% | 11.73% | 1.692% | 5.956% |  | 3.721%
| PDM3 |  | -44.36% | -15.14% | -18.36% | -15.92% | 43.71% | -99.09% | 2.31% | 52.05% |  | 4.147%
| Z-Score (Absolute) |  | 1.16 | 0.05 | 0.08 | 0 | 2 | 2.8 | 0.61 | 2.28 | NA | 0.67
| Analytical results for Hf in OREAS 465 (Certified Value 14.4 ppm) | [15 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 1
1
| 1 | 1 | 14.8 | 14.2 | 14.3 | 12 | 15.4 | 14.4271 | 14.5 | 0.74 | 11.41 | 40 |  |  |  |  | 1
| 1 | 2 | 16.7 | 13.3 | 13.8 | 10.9 | 14.5 | 15.4753 | 12 | 0.73 | 16.56 | 42 |  |  |  |  | 9
| 1 | 3 | 13.5 | 14.1 | 14.4 | 12.8 | 13.9 | 14.3944 | 9.5 | 0.76 | 15.43 | 42 |  |  |  |  | 16
| 1 | 4 | 17.6 | 13.6 | 14.7 | 12.8 | 15 | 12.2891 | 16.9 | 0.82 | 11.32 | 42 |  |  |  |  | 14.4049
| 1 | 5 | 14.3 | 14.5 | 14.6 | 15.6 | 15.6 | 14.0762 | 16.9 | 0.63 | 20.01 | 42 |  |  |  |  | 79
| 1 | 6 | 15 | 13.7 | 12.7 | 13.8 | 14.7 | 12.9428 | 17.4 | 0.69 | 17.44 | 40
| Mean |  | 15.3167 | 13.9 | 14.0833 | 12.9833 | 14.85 | 13.9341 | 14.5333 | 0.72833 | 15.3617 | 41.3333
| Median |  | 14.9 | 13.9 | 14.35 | 12.8 | 14.85 | 14.2353 | 15.7 | 0.735 | 15.995 | 42
| Std Dev. |  | 1.5381 | 0.44272 | 0.74677 | 1.6055 | 0.62209 | 1.1436 | 3.1992 | 0.064317 | 3.4444 | 1.0328
| Rel.Std.Dev. |  | 10.04% | 3.185% | 5.303% | 12.37% | 4.189% | 8.207% | 22.01% | 8.831% | 22.42% | 2.499%
| PDM3 |  | 6.33% | -3.505% | -2.232% | -9.869% | 3.09% | -3.268% | 0.8916% | -94.94% | 6.642% | 186.9%
| Z-Score (Absolute) |  | 0.83 | 0.52 | 0.08 | 1.39 | 0.38 | 0.49 | 0.08 | 13.06 | 0.87 | 25.6
| Analytical results for Ho in OREAS 465 (Certified Value 26.8 ppm) | [15 empty] | Certified
Ho | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 29.4 | 25.3 | 27.4 | 25.8 | 27.2 | 25.8216 | 26.1 | 26.4 | 19.4 | 26.69 | 31.0545 | 26.6 | 28.2 |  | 1
| 1 | 2 | 28 | 23.3 | 25.7 | 25.6 | 27.6 | 26.6212 | 26.3 | 26.8 | 19.2 | 26.37 | 31.5841 | 26.6 | 28.3 |  | 10
| 1 | 3 | 27.5 | 25.9 | 26.3 | 27.1 | 26.6 | 26.6508 | 25.9 | 26.3 | 19.2 | 27.23 | 32.4822 | 26.55 | 27.1 |  | 16
| 1 | 4 | 27.8 | 24.3 | 28.1 | 26.9 | 26.6 | 25.0343 | 25.9 | 27.3 | 19.6 | 26.95 | 31.2911 | 26.45 | 27.6 |  | 26.7772
| 1 | 5 | 27.3 | 25.1 | 28.2 | 26.7 | 28.9 | 25.6445 | 26.5 | 26.4 | 20.2 | 26.91 | 30.6668 | 26.9 | 27.5 |  | 80
| 1 | 6 | 26.1 | 28.1 | 27.3 | 27.1 | 27.8 | 27.032 | 26.6 | 25.9 | 19.5 | 26.35 | 30.6346 | 26.55 | 27.8
| Mean |  | 27.6833 | 25.3333 | 27.1667 | 26.5333 | 27.45 | 26.1341 | 26.2167 | 26.5167 | 19.5167 | 26.75 | 31.2856 | 26.6083 | 27.75
| Median |  | 27.65 | 25.2 | 27.35 | 26.8 | 27.4 | 26.2214 | 26.2 | 26.4 | 19.45 | 26.8 | 31.1728 | 26.575 | 27.7
| Std Dev. |  | 1.0722 | 1.6269 | 0.9913 | 0.66533 | 0.8666 | 0.75599 | 0.29944 | 0.47924 | 0.37103 | 0.34756 | 0.69007 | 0.15303 | 0.45056
| Rel.Std.Dev. |  | 3.873% | 6.422% | 3.649% | 2.508% | 3.157% | 2.893% | 1.142% | 1.807% | 1.901% | 1.299% | 2.206% | 0.5751% | 1.624%
| PDM3 |  | 3.384% | -5.392% | 1.455% | -0.9107% | 2.513% | -2.402% | -2.093% | -0.9729% | -27.11% | -0.1015% | 16.84% | -0.6306% | 3.633%
| Z-Score (Absolute) |  | 1.3 | 1.05 | 0.67 | 0.09 | 1.02 | 0.57 | 0.47 | 0.11 | 8.56 | 0.17 | 5.65 | 0 | 1.38
| Analytical results for In in OREAS 465 (Certified Value 3.18 ppm) | [15 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  | 2
3
| 1 | 1 | 3.33 | 3.03 | 3.71 | 3.2 | 3.35 | 3.1265 | 3.2 | 2.97 | 3.24 | 3.07 | 3.1 | 2.76 |  |  | 1
| 1 | 2 | 3.5 | 2.93 | 3.7 | 3.19 | 3.44 | 3.0726 | 3.4 | 3.09 | 3.17 | 3.06 | 3.1 | 2.78 |  |  | 0
| 1 | 3 | 3.18 | 3 | 3.74 | 3.34 | 3.23 | 2.9876 | 3.4 | 3.11 | 3.17 | 3.07 | 3 | 2.82 |  |  | 16
| 1 | 4 | 3.47 | 3.03 | 3.5 | 3.41 | 3.3 | 2.989 | 3.4 | 3.21 | 3.21 | 3.22 | 3 | 2.88 |  |  | 3.1831
| 1 | 5 | 3.31 | 3.16 | 3.56 | 3.05 | 3.19 | 3.047 | 3.3 | 3.39 | 3.19 | 3.25 | 3.1 | 2.94 |  |  | 81
| 1 | 6 | 3.35 | 3.05 | 3.27 | 3.2 | 3.16 | 3.0721 | 3.4 | 3.07 | 3.22 | 3.18 | 3 | 2.84
| Mean |  | 3.3567 | 3.0333 | 3.58 | 3.2317 | 3.2783 | 3.0491 | 3.35 | 3.14 | 3.2 | 3.1417 | 3.05 | 2.8367
| Median |  | 3.34 | 3.03 | 3.63 | 3.2 | 3.265 | 3.0595 | 3.4 | 3.1 | 3.2 | 3.125 | 3.05 | 2.83
| Std Dev. |  | 0.11622 | 0.075011 | 0.17877 | 0.12671 | 0.10572 | 0.053802 | 0.083666 | 0.14464 | 0.028284 | 0.085186 | 0.054772 | 0.066232
| Rel.Std.Dev. |  | 3.462% | 2.473% | 4.994% | 3.921% | 3.225% | 1.765% | 2.497% | 4.606% | 0.8839% | 2.711% | 1.796% | 2.335%
| PDM3 |  | 5.452% | -4.706% | 12.47% | 1.525% | 2.991% | -4.21% | 5.243% | -1.355% | 0.5302% | -1.302% | -4.182% | -10.88%
| Z-Score (Absolute) |  | 1.03 | 0.76 | 2.28 | 0.34 | 0.6 | 0.68 | 1 | 0.45 | 0.16 | 0.16 | 0.67 | 1.86
| Analytical results for K in OREAS 465 (Indicative Value 0.023 wt.%) | [15 empty] | Indicative
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 3
3
| 1 | 1 | 0.01 | 0.03 | 0.03 | < 0.01 | 0.04 | 0.11 | < 0.02 | 0.017 | < 0.002 | 0.01 | 0.01 | < 0.01 |  |  | 1
| 1 | 2 | 0.01 | 0.1 | 0.03 | < 0.01 | 0.04 | 0.11 | < 0.02 | 0.017 | < 0.002 | 0.01 | 0.01 | < 0.01 |  |  | 11
| 1 | 3 | < 0.01 | 0.03 | 0.03 | < 0.01 | 0.04 | 0.11 | 0.02 | 0.016 | < 0.002 | 0.01 | < 0.01 | < 0.01 |  |  | 16
| 1 | 4 | < 0.01 | 0.03 | 0.07 | < 0.01 | 0.04 | 0.11 | < 0.02 | 0.017 | < 0.002 | 0.01 | < 0.01 | < 0.01 |  |  | 0.023204
| 1 | 5 | < 0.01 | 0.04 | 0.07 | 0.02 | 0.04 | 0.11 | < 0.02 | 0.017 | < 0.002 | 0.01 | < 0.01 | < 0.01 |  |  | 32
| 1 | 6 | < 0.01 | 0.03 | 0.07 | 0.02 | 0.04 | 0.11 | < 0.02 | 0.017 | < 0.002 | 0.01 | < 0.01 | < 0.01
| Mean |  | 0.01 | 0.043333 | 0.05 | 0.02 | 0.04 | 0.11 | 0.02 | 0.016833 |  | 0.01 | 0.01
| Median |  | 0.01 | 0.03 | 0.05 | 0.02 | 0.04 | 0.11 | 0.02 | 0.017 |  | 0.01 | 0.01
| Std Dev. |  | 0 | 0.028048 | 0.021909 | 0 | 0 | 0 |  | 0.00040825 |  | 0 | 0
| Rel.Std.Dev. |  | 0% | 64.73% | 43.82% | 0% | 0% | 0% |  | 2.425% |  | 0% | 0%
| PDM3 |  | -56.9% | 86.75% | 115.5% | -13.81% | 72.39% | 374.1% | -13.81% | -27.45% |  | -56.9% | -56.9%
| Z-Score (Absolute) |  | 0.67 | 0.81 | 2.02 | 0 | 1.35 | 6.07 | 0 | 0.21 | NA | 0.67 | 0.67 | NA
| Analytical results for La in OREAS 465 (Certified Value 2.27 wt.%) | [15 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | > 1 | > 1 | > 1 | > 1 | > 1 | 2.4882 | 2.19 | 2.201 | > 0.5 | > 0.5 | 2.4233 | 2.1864 | 2.26 |  | 1
| 1 | 2 | > 1 | > 1 | > 1 | > 1 | > 1 | 2.5121 | 2.2088 | 2.1957 | > 0.5 | > 0.5 | 2.3964 | 2.2146 | 2.27 |  | 0
| 1 | 3 | > 1 | > 1 | > 1 | > 1 | > 1 | 2.5157 | 2.1838 | 2.1851 | > 0.5 | > 0.5 | 2.4968 | 2.2387 | 2.31 |  | 16
| 1 | 4 | > 1 | > 1 | > 1 | > 1 | > 1 | 2.5061 | 2.2181 | 2.3083 | > 0.5 | > 0.5 | 2.4522 | 2.2199 | 2.34 |  | 2.2739
| 1 | 5 | > 1 | > 1 | > 1 | > 1 | > 1 | 2.5077 | 2.2259 | 2.2354 | > 0.5 | > 0.5 | 2.3916 | 2.2667 | 2.37 |  | 82
| 1 | 6 | > 1 | > 1 | > 1 | > 1 | > 1 | 2.5308 | 2.2328 | 2.1863 | > 0.5 | > 0.5 | 2.4232 | 2.2342 | 2.26
| Mean |  | 07: 2.5101; 08: 2.2099; 09: 2.2186; 12: 2.4306; 17: 2.2267; 22: 2.3017
| Median |  | 07: 2.5099; 08: 2.2134; 09: 2.1983; 12: 2.4232; 17: 2.227; 22: 2.29
| Std Dev. |  | 07: 0.013913; 08: 0.019645; 09: 0.047614; 12: 0.039108; 17: 0.026904; 22: 0.046224
| Rel.Std.Dev. |  | 07: 0.5543%; 08: 0.8889%; 09: 2.146%; 12: 1.609%; 17: 1.208%; 22: 2.008%
| PDM3 |  | 07: 10.39%; 08: -2.816%; 09: -2.431%; 12: 6.889%; 17: -2.074%; 22: 1.221%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 2.81 | 0.62 | 0.73 | NA | NA | 1.9 | 0.43 | 0.43
| Analytical results for Li in OREAS 465 (Certified Value 3.04 ppm) | [15 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 2
2
| 1 | 1 | 2.9 | 2.5 | 3 | 3.1 | 1.7 | 3.4756 | 2.8 | 3 | 3.2 | < 0.1 | 6.1505 | 5 | < 5 | 3 | 1
| 1 | 2 | 3 | 2.9 | 2.9 | 3.2 | 1.6 | 3.6465 | 2.9 | 3 | 3.2 | < 0.1 | 6.1116 | 5 | < 5 | 3 | 16
| 1 | 3 | 2.8 | 2.3 | 2.8 | 3.3 | 1.5 | 3.3934 | 2.8 | 3 | 3.2 | < 0.1 | 6.1761 | 5 | < 5 | 3 | 16
| 1 | 4 | 3 | 2.4 | 3.2 | 3.3 | 1.5 | 3.5981 | 2.9 | 4 | 3.3 | < 0.1 | 6.3057 | 5 | < 5 | 3 | 3.0419
| 1 | 5 | 3 | 2.4 | 3.3 | 4.5 | 2.8 | 3.6644 | 2.8 | 3 | 3.3 | < 0.1 | 6.2631 | 5 | < 5 | 3 | 83
| 1 | 6 | 3.4 | 2.5 | 2.8 | 4.6 | 1.8 | 3.4721 | 2.9 | 3 | 3.2 | < 0.1 | 6.2304 | 5 | < 5 | 3
| Mean |  | 3.0167 | 2.5 | 3 | 3.6667 | 1.8167 | 3.5417 | 2.85 | 3.1667 | 3.2333 |  | 6.2062 | 5 |  | 3
| Median |  | 3 | 2.45 | 2.95 | 3.3 | 1.65 | 3.5369 | 2.85 | 3 | 3.2 |  | 6.2033 | 5 |  | 3
| Std Dev. |  | 0.20412 | 0.20976 | 0.20976 | 0.68896 | 0.49565 | 0.10992 | 0.054772 | 0.40825 | 0.05164 |  | 0.073072 | 0 |  | 0
| Rel.Std.Dev. |  | 6.767% | 8.39% | 6.992% | 18.79% | 27.28% | 3.104% | 1.922% | 12.89% | 1.597% |  | 1.177% | 0% |  | 0%
| PDM3 |  | -0.8281% | -17.81% | -1.376% | 20.54% | -40.28% | 16.43% | -6.307% | 4.103% | 6.295% |  | 104% | 64.37% |  | -1.376%
| Z-Score (Absolute) |  | 0.5 | 2.33 | 0.29 | 0.5 | 5.15 | 1.61 | 0.82 | 0.29 | 0.53 | NA | 10.99 | 6.74 | NA | 0.29
| Analytical results for Lu in OREAS 465 (Certified Value 1.72 ppm) | [15 empty] | Certified
Lu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 3.24 | 2.01 | 1.98 | 1.79 | 2.09 | 1.2994 | 1.5 | 1.39 | 1.51 | 1.48 | 2.1905 | 1.5 | 1.66 |  | 1
| 1 | 2 | 3.21 | 1.83 | 1.91 | 1.8 | 2.17 | 1.2329 | 1.6 | 1.37 | 1.43 | 1.52 | 2.1555 | 1.7 | 1.64 |  | 1
| 1 | 3 | 2.99 | 1.97 | 1.9 | 1.88 | 2.1 | 1.2923 | 1.6 | 1.47 | 1.41 | 1.44 | 2.2212 | 1.6 | 1.68 |  | 16
| 1 | 4 | 3.22 | 1.95 | 2.02 | 1.83 | 2.08 | 1.2411 | 1.6 | 1.43 | 1.45 | 1.58 | 2.1143 | 1.7 | 1.72 |  | 1.7206
| 1 | 5 | 2.96 | 2 | 2.16 | 1.94 | 2.13 | 1.3144 | 1.7 | 1.39 | 1.43 | 1.74 | 2.0782 | 1.5 | 1.68 |  | 84
| 1 | 6 | 3.84 | 2.01 | 1.91 | 1.93 | 2.09 | 1.3336 | 1.7 | 1.29 | 1.45 | 1.52 | 1.8569 | 1.7 | 1.7
| Mean |  | 3.2433 | 1.9617 | 1.98 | 1.8617 | 2.11 | 1.2856 | 1.6167 | 1.39 | 1.4467 | 1.5467 | 2.1028 | 1.6167 | 1.68
| Median |  | 3.215 | 1.985 | 1.945 | 1.855 | 2.095 | 1.2958 | 1.6 | 1.39 | 1.44 | 1.52 | 2.1349 | 1.65 | 1.68
| Std Dev. |  | 0.3169 | 0.068823 | 0.1002 | 0.064936 | 0.034059 | 0.040337 | 0.075277 | 0.060663 | 0.034448 | 0.10558 | 0.13091 | 0.098319 | 0.028284
| Rel.Std.Dev. |  | 9.771% | 3.508% | 5.061% | 3.488% | 1.614% | 3.138% | 4.656% | 4.364% | 2.381% | 6.826% | 6.225% | 6.082% | 1.684%
| PDM3 |  | 88.5% | 14.01% | 15.07% | 8.197% | 22.63% | -25.28% | -6.042% | -19.22% | -15.92% | -10.11% | 22.21% | -6.042% | -2.361%
| Z-Score (Absolute) |  | 3.46 | 0.67 | 0.72 | 0.43 | 1.03 | 0.94 | 0.15 | 0.69 | 0.56 | 0.32 | 1.13 | 0.15 | 0
| Analytical results for Mg in OREAS 465 (Certified Value 0.374 wt.%) | [15 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.35 | 0.35 | 0.19 | 0.21 | 0.17 | 0.40101 | 0.37 | 0.38 | 0.395 | 0.3783 | 0.37206 | 0.33 | 0.398 | 0.38 | 1
| 1 | 2 | 0.35 | 0.31 | 0.19 | 0.2 | 0.18 | 0.4051 | 0.37 | 0.37 | 0.397 | 0.3974 | 0.37099 | 0.33 | 0.3995 | 0.38 | 2
| 1 | 3 | 0.32 | 0.35 | 0.18 | 0.21 | 0.18 | 0.41074 | 0.37 | 0.38 | 0.399 | 0.3902 | 0.37942 | 0.32 | 0.4 | 0.38 | 16
| 1 | 4 | 0.35 | 0.36 | 0.19 | 0.2 | 0.18 | 0.4101 | 0.37 | 0.4 | 0.395 | 0.3726 | 0.37092 | 0.33 | 0.402 | 0.39 | 0.37383
| 1 | 5 | 0.32 | 0.35 | 0.19 | 0.34 | 0.18 | 0.406 | 0.37 | 0.41 | 0.399 | 0.3624 | 0.37782 | 0.32 | 0.402 | 0.38 | 85
| 1 | 6 | 0.34 | 0.34 | 0.19 | 0.36 | 0.17 | 0.4071 | 0.37 | 0.38 | 0.396 | 0.3827 | 0.37176 | 0.33 | 0.4015 | 0.41
| Mean |  | 0.33833 | 0.34333 | 0.18833 | 0.25333 | 0.17667 | 0.40667 | 0.37 | 0.38667 | 0.39683 | 0.3806 | 0.37383 | 0.32667 | 0.4005 | 0.38667
| Median |  | 0.345 | 0.35 | 0.19 | 0.21 | 0.18 | 0.40655 | 0.37 | 0.38 | 0.3965 | 0.3805 | 0.37191 | 0.33 | 0.40075 | 0.38
| Std Dev. |  | 0.01472 | 0.017512 | 0.0040825 | 0.075277 | 0.005164 | 0.0035645 | 6.0809e-17 | 0.015055 | 0.0018348 | 0.012485 | 0.0037692 | 0.005164 | 0.0016125 | 0.012111
| Rel.Std.Dev. |  | 4.351% | 5.101% | 2.168% | 29.71% | 2.923% | 0.8765% | 1.643e-14% | 3.894% | 0.4624% | 3.28% | 1.008% | 1.581% | 0.4026% | 3.132%
| PDM3 |  | -9.495% | -8.157% | -49.62% | -32.23% | -52.74% | 8.787% | -1.024% | 3.435% | 6.154% | 1.812% | 1.544e-05% | -12.62% | 7.135% | 3.435%
| Z-Score (Absolute) |  | 0.85 | 0.55 | 4.63 | 3.53 | 4.92 | 0.88 | 0.05 | 0.37 | 0.63 | 0.22 | 0.05 | 1.14 | 0.72 | 0.25
| Analytical results for Mn in OREAS 465 (Certified Value 0.198 wt.%) | [15 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.178 | 0.188 | 0.0841 | 0.0936 | 0.0763 | 0.204 | 0.20078 | 0.1832 | 0.193 | 0.2113 | 0.1561 | 0.2075 | 0.25 |  | 1
| 1 | 2 | 0.181 | 0.172 | 0.0857 | 0.0902 | 0.0773 | 0.20668 | 0.20515 | 0.1767 | 0.194 | 0.2056 | 0.1561 | 0.2095 | 0.242 |  | 3
| 1 | 3 | 0.17 | 0.188 | 0.0836 | 0.0952 | 0.0811 | 0.20788 | 0.20154 | 0.1899 | 0.193 | 0.2082 | 0.1546 | 0.2095 | 0.246 |  | 16
| 1 | 4 | 0.179 | 0.192 | 0.0891 | 0.0921 | 0.0781 | 0.20571 | 0.20509 | 0.1992 | 0.192 | 0.2132 | 0.1548 | 0.208 | 0.245 |  | 0.19771
| 1 | 5 | 0.172 | 0.193 | 0.0888 | 0.164 | 0.0793 | 0.20158 | 0.19906 | 0.2094 | 0.193 | 0.21 | 0.1558 | 0.2085 | 0.245 |  | 86
| 1 | 6 | 0.174 | 0.19 | 0.0889 | 0.17 | 0.0759 | 0.20251 | 0.20579 | 0.1863 | 0.192 | 0.2145 | 0.1568 | 0.2065 | 0.256
| Mean |  | 0.17567 | 0.18717 | 0.0867 | 0.11752 | 0.078 | 0.20473 | 0.2029 | 0.19078 | 0.19283 | 0.21047 | 0.1557 | 0.20825 | 0.24733
| Median |  | 0.176 | 0.189 | 0.08725 | 0.0944 | 0.0777 | 0.20486 | 0.20331 | 0.1881 | 0.193 | 0.21065 | 0.15595 | 0.20825 | 0.2455
| Std Dev. |  | 0.0043205 | 0.007705 | 0.0025448 | 0.038412 | 0.0019545 | 0.0024543 | 0.0028036 | 0.011783 | 0.00075277 | 0.0032703 | 0.0008438 | 0.0011726 | 0.0049666
| Rel.Std.Dev. |  | 2.459% | 4.117% | 2.935% | 32.69% | 2.506% | 1.199% | 1.382% | 6.176% | 0.3904% | 1.554% | 0.5419% | 0.5631% | 2.008%
| PDM3 |  | -11.15% | -5.334% | -56.15% | -40.56% | -60.55% | 3.548% | 2.624% | -3.505% | -2.468% | 6.451% | -21.25% | 5.329% | 25.1%
| Z-Score (Absolute) |  | 0.58 | 0.02 | 4.02 | 2.83 | 4.35 | 0.54 | 0.47 | 0 | 0.08 | 0.76 | 1.35 | 0.67 | 2.12
| Analytical results for Mo in OREAS 465 (Certified Value 98 ppm) | [15 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  | 0
0
| 1 | 1 | 88.9 | 88.3 | 98 | 92.7 | 96.1 | 109.4153 | 106.8 | 107 | 112 | 104.6 | 85.7899 | 81.5 |  |  | 1
| 1 | 2 | 90.2 | 86.2 | 97.1 | 93.1 | 99.5 | 107.0979 | 109 | 106 | 113 | 104.6 | 85.9653 | 82 |  |  | 0
| 1 | 3 | 84.6 | 90 | 97.1 | 98.8 | 94.7 | 104.4822 | 106.3 | 105 | 114 | 105.7 | 85.3297 | 83.5 |  |  | 16
| 1 | 4 | 90.2 | 90.2 | 98.5 | 96.7 | 94.8 | 108.4001 | 108.5 | 111 | 114 | 108.4 | 86.0589 | 85.5 |  |  | 98.0101
| 1 | 5 | 86.1 | 88.2 | 99.3 | 91.4 | 92.5 | 103.3617 | 107.1 | 115 | 115 | 109.4 | 86.626 | 87 |  |  | 87
| 1 | 6 | 84.1 | 87.7 | 91.5 | 94.5 | 92.1 | 104.7336 | 109.8 | 109 | 112 | 105.9 | 86.2646 | 85
| Mean |  | 87.35 | 88.4333 | 96.9167 | 94.5333 | 94.95 | 106.2485 | 107.9167 | 108.8333 | 113.3333 | 106.4333 | 86.0057 | 84.0833
| Median |  | 87.5 | 88.25 | 97.55 | 93.8 | 94.75 | 105.9157 | 107.8 | 108 | 113.5 | 105.8 | 86.0121 | 84.25
| Std Dev. |  | 2.7689 | 1.4949 | 2.7845 | 2.7616 | 2.692 | 2.4136 | 1.3848 | 3.7103 | 1.2111 | 2.0106 | 0.43814 | 2.1311
| Rel.Std.Dev. |  | 3.17% | 1.69% | 2.873% | 2.921% | 2.835% | 2.272% | 1.283% | 3.409% | 1.069% | 1.889% | 0.5094% | 2.535%
| PDM3 |  | -10.88% | -9.771% | -1.116% | -3.547% | -3.122% | 8.406% | 10.11% | 11.04% | 15.63% | 8.594% | -12.25% | -14.21%
| Z-Score (Absolute) |  | 0.62 | 0.55 | 0.1 | 0.13 | 0.1 | 0.67 | 0.78 | 0.84 | 1.15 | 0.68 | 0.72 | 0.85
| Analytical results for Na in OREAS 465 (Certified Value < 0.2 wt.%) | [15 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 3
3
| 1 | 1 | 0.09 | < 0.01 | 0.37 | 0.32 | 0.3 | 0.08 | 0.08 | 0.084 | 0.0453 | 0.123 | 0.1 | 0.07 | 0.08 |  | 1
| 1 | 2 | 0.08 | < 0.01 | 0.36 | 0.32 | 0.31 | 0.08 | 0.08 | 0.086 | 0.0442 | 0.12551 | 0.1 | 0.07 | 0.08 |  | 4
| 1 | 3 | 0.08 | < 0.01 | 0.36 | 0.33 | 0.31 | 0.08 | 0.08 | 0.086 | 0.0453 | 0.12572 | 0.09 | 0.07 | 0.08 |  | 16
| 1 | 4 | 0.09 | < 0.01 | 0.06 | 0.32 | 0.32 | 0.08 | 0.08 | 0.085 | 0.0443 | 0.12242 | 0.1 | 0.07 | 0.08 |  | < 0.2
| 1 | 5 | 0.09 | < 0.01 | 0.06 | 0.07 | 0.31 | 0.08 | 0.08 | 0.084 | 0.0436 | 0.12279 | 0.1 | 0.07 | 0.07 |  | 88
| 1 | 6 | 0.12 | < 0.01 | 0.06 | 0.09 | 0.29 | 0.08 | 0.08 | 0.083 | 0.0452 | 0.12501 | 0.1 | 0.07 | 0.08
| Mean |  | 0.091667 |  | 0.21167 | 0.24167 | 0.30667 | 0.08 | 0.08 | 0.084667 | 0.04465 | 0.12408 | 0.098333 | 0.07 | 0.078333
| Median |  | 0.09 |  | 0.21 | 0.32 | 0.31 | 0.08 | 0.08 | 0.0845 | 0.04475 | 0.12401 | 0.1 | 0.07 | 0.08
| Std Dev. |  | 0.01472 |  | 0.16618 | 0.12545 | 0.010328 | 0 | 0 | 0.0012111 | 0.00071764 | 0.0014955 | 0.0040825 | 0 | 0.0040825
| Rel.Std.Dev. |  | 16.06% |  | 78.51% | 51.91% | 3.368% | 0% | 0% | 1.43% | 1.607% | 1.205% | 4.152% | 0% | 5.212%
PDM3
| Z-Score (Absolute) |  | 0.03 | NA | 6.01 | 9.08 | 10.54 | 0.25 | 0.25 | 0.03 | 1.94 | 1.84 | 0.62 | 0.73 | 0.33
| Analytical results for Nb in OREAS 465 (Indicative Value 399 ppm) | [15 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 338 | 480 | 480 | 359 | 230 | 4022.2861 | 3141.7 | 818.6 | > 500 | FL |  |  |  |  | 1
| 1 | 2 | 394 | 420 | 378 | 273 | 230 | 4183.6518 | 3275.7 | 247.9 | > 500 | FL |  |  |  |  | 27
| 1 | 3 | 291 | > 500 | 408 | 417 | 250 | 3908.8851 | 3239 | 180.6 | > 500 | FL |  |  |  |  | 16
| 1 | 4 | 460 | > 500 | 420 | 362 | 20 | 4238.2069 | 3351.3 | 825.3 | > 500 | FL |  |  |  |  | 398.9111
| 1 | 5 | 329 | > 500 | 372 | > 500 | 290 | 4106.8087 | 3285 | 480 | > 500 | FL |  |  |  |  | 33
| 1 | 6 | 400 | 490 | 351 | > 500 | 220 | 3903.7742 | 3243.3 | 826.9 | > 500 | FL
| Mean |  | 368.6667 | 463.3333 | 401.5 | 352.75 | 206.6667 | 4060.6021 | 3256 | 563.2167
| Median |  | 366 | 480 | 393 | 360.5 | 230 | 4064.5474 | 3259.5 | 649.3
| Std Dev. |  | 60.8463 | 37.8594 | 45.8683 | 59.4776 | 94.798 | 139.9292 | 69.0304 | 302.0538
| Rel.Std.Dev. |  | 16.5% | 8.171% | 11.42% | 16.86% | 45.87% | 3.446% | 2.12% | 53.63%
| PDM3 |  | -7.582% | 16.15% | 0.649% | -11.57% | -48.19% | 917.9% | 716.2% | 41.19%
| Z-Score (Absolute) |  | 0.41 | 0.2 | 0.2 | 0.51 | 1.21 | 23.25 | 18.09 | 0.84 | NA | NA
| Analytical results for Nd in OREAS 465 (Certified Value 1.1 wt.%) | [15 empty] | Certified
Nd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | > 1 | > 1 | > 1 | > 1 | > 1 | 1.2007 | 1.0637 | 1.063 | > 0.5 | 1.1333 | 1.0489 | 1.0575 | 1.12 |  | 1
| 1 | 2 | > 1 | > 1 | > 1 | > 1 | > 1 | 1.226 | 1.086 | 1.0413 | > 0.5 | 1.1331 | 1.0212 | 1.0941 | 1.11 |  | 0
| 1 | 3 | > 1 | > 1 | > 1 | > 1 | > 1 | 1.27 | 1.0602 | 1.0588 | > 0.5 | 1.155 | 1.0651 | 1.1016 | 1.12 |  | 16
| 1 | 4 | > 1 | > 1 | > 1 | > 1 | > 1 | 1.2274 | 1.0751 | 1.1268 | > 0.5 | 1.1826 | 1.0078 | 1.0889 | 1.15 |  | 1.0951
| 1 | 5 | > 1 | > 1 | > 1 | > 1 | > 1 | 1.2612 | 1.1037 | 1.1233 | > 0.5 | 1.1481 | 1.0296 | 1.1251 | 1.18 |  | 89
| 1 | 6 | > 1 | > 1 | > 1 | > 1 | > 1 | 1.2239 | 1.0897 | 1.0782 | > 0.5 | 1.1667 | 1.0438 | 1.1093 | 1.12
| Mean |  | 07: 1.2349; 08: 1.0797; 09: 1.0819; 11: 1.1531; 12: 1.0361; 17: 1.0961; 22: 1.1333
| Median |  | 07: 1.2267; 08: 1.0806; 09: 1.0706; 11: 1.1515; 12: 1.0367; 17: 1.0978; 22: 1.12
| Std Dev. |  | 07: 0.025874; 08: 0.016555; 09: 0.035448; 11: 0.019352; 12: 0.020645; 17: 0.022768; 22: 0.026583
| Rel.Std.Dev. |  | 07: 2.095%; 08: 1.533%; 09: 3.276%; 11: 1.678%; 12: 1.993%; 17: 2.077%; 22: 2.346%
| PDM3 |  | 07: 12.76%; 08: -1.407%; 09: -1.21%; 11: 5.294%; 12: -5.396%; 17: 0.08451%; 22: 3.487%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 3.35 | 0.39 | 0.34 | NA | 1.38 | 1.45 | 0 | 0.67
| Analytical results for Ni in OREAS 465 (Certified Value 106 ppm) | [15 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES |  |  |  | 0
0
| 1 | 1 | 447 | 107.5 | 112.5 | 243 | 228 | 101.8841 | 86 | 111 | 103 | 80 | 122 |  |  |  | 1
| 1 | 2 | 458 | 104.5 | 108 | 233 | 232 | 117.1332 | 87 | 113 | 97 | 80 | 118 |  |  |  | 6
| 1 | 3 | 430 | 100 | 104.5 | 245 | 235 | 104.666 | 79 | 117 | 102 | 85 | 130 |  |  |  | 16
| 1 | 4 | 457 | 107.5 | 112.5 | 237 | 237 | 109.5719 | 114 | 115 | 106 | 85 | 126 |  |  |  | 105.8075
| 1 | 5 | 434 | 103 | 111.5 | 430 | 236 | 114.0994 | 104 | 115 | 112 | 90 | 122 |  |  |  | 90
| 1 | 6 | 406 | 106 | 104.5 | 436 | 230 | 114.4042 | 82 | 116 | 123 | 85 | 130
| Mean |  | 438.6667 | 104.75 | 108.9167 | 304 | 233 | 110.2931 | 92 | 114.5 | 107.1667 | 84.1667 | 124.6667
| Median |  | 440.5 | 105.25 | 109.75 | 244 | 233.5 | 111.8357 | 86.5 | 115 | 104.5 | 85 | 124
| Std Dev. |  | 19.7146 | 2.9112 | 3.8002 | 100.032 | 3.5777 | 6.0179 | 13.842 | 2.1679 | 9.196 | 3.7639 | 4.8442
| Rel.Std.Dev. |  | 4.494% | 2.779% | 3.489% | 32.91% | 1.535% | 5.456% | 15.05% | 1.893% | 8.581% | 4.472% | 3.886%
| PDM3 |  | 314.6% | -0.9994% | 2.939% | 187.3% | 120.2% | 4.239% | -13.05% | 8.215% | 1.285% | -20.45% | 17.82%
| Z-Score (Absolute) |  | 15.41 | 0.26 | 0.06 | 9.09 | 5.76 | 0 | 0.86 | 0.2 | 0.15 | 1.23 | 0.67
| Analytical results for P in OREAS 465 (Certified Value 3.15 wt.%) | [15 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 2
2
| 1 | 1 | > 1 | > 1 | > 1 | > 1 | > 1 | 3.551 | 3.66 | 3.55 | 3.2242 | 2.8904 | 2.7 | 3.1195 | 2.84 |  | 1
| 1 | 2 | > 1 | > 1 | > 1 | > 1 | > 1 | 3.605 | 3.55 | 3.52 | 3.0043 | 2.8436 | 2.7 | 3.0415 | 2.86 |  | 7
| 1 | 3 | > 1 | > 1 | > 1 | > 1 | > 1 | 3.574 | 3.49 | 3.51 | 3.1132 | 2.8797 | 2.7 | 2.7165 | 2.72 |  | 16
| 1 | 4 | > 1 | > 1 | > 1 | > 1 | > 1 | 3.537 | 3.69 | 3.59 | 3.2915 | 3.0265 | 2.6 | 2.6905 | 3.01 |  | 3.1508
| 1 | 5 | > 1 | > 1 | > 1 | > 1 | > 1 | 3.579 | 3.74 | 3.54 | 3.0884 | 3.0073 | 2.7 | 2.654 | 2.74 |  | 91
| 1 | 6 | > 1 | > 1 | > 1 | > 1 | > 1 | 3.573 | 3.76 | 3.52 | 3.2331 | 2.8222 | 2.7 | 2.804 | 2.98
| Mean | [6 empty] | 3.5698 | 3.6483 | 3.5383 | 3.1591 | 2.9116 | 2.6833 | 2.8377 | 2.8583
| Median | [6 empty] | 3.5735 | 3.675 | 3.53 | 3.1687 | 2.885 | 2.7 | 2.7603 | 2.85
| Std Dev. | [6 empty] | 0.023583 | 0.10722 | 0.029269 | 0.10799 | 0.085359 | 0.040825 | 0.19605 | 0.1194
| Rel.Std.Dev. | [6 empty] | 0.6606% | 2.939% | 0.8272% | 3.418% | 2.932% | 1.521% | 6.909% | 4.177%
| PDM3 | [6 empty] | 13.3% | 15.79% | 12.3% | 0.2633% | -7.592% | -14.84% | -9.939% | -9.283%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 1.31 | 1.5 | 1.23 | 0.3 | 0.3 | 0.86 | 0.49 | 0.43
| Analytical results for Pb in OREAS 465 (Certified Value 573 ppm) | [15 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 564 | 580 | 580 | 612 | 574 | 551.9377 | 546 | 572 | 480.9 | 577.585 | 620 | 497 |  |  | 1
| 1 | 2 | 571 | 542 | 559 | 573 | 595 | 561.1046 | 549 | 581 | 484.2 | 579.745 | 630 | 493 |  |  | 8
| 1 | 3 | 561 | 576 | 566 | 599 | 593 | 550.4525 | 552 | 585 | 505.4 | 580.717 | 630 | 501 |  |  | 16
| 1 | 4 | 569 | 587 | 590 | 596 | 590 | 557.1248 | 537 | 592 | 510 | 579.305 | 630 | 514 |  |  | 572.5402
| 1 | 5 | 561 | 580 | 580 | 550 | 571 | 557.4403 | 543 | 594 | 480.6 | 579.108 | 620 | 519 |  |  | 92
| 1 | 6 | 565 | 569 | 588 | 575 | 575 | 563.9576 | 551 | 569 | 484.7 | 580.294 | 630 | 512
| Mean |  | 565.1667 | 572.3333 | 577.1667 | 584.1667 | 583 | 557.0029 | 546.3333 | 582.1667 | 490.9667 | 579.459 | 626.6667 | 506
| Median |  | 564.5 | 578 | 580 | 585.5 | 582.5 | 557.2825 | 547.5 | 583 | 484.45 | 579.525 | 630 | 506.5
| Std Dev. |  | 4.1191 | 15.9833 | 12.2706 | 22.4091 | 10.7889 | 5.1761 | 5.6451 | 10.2258 | 13.1488 | 1.0973 | 5.164 | 10.4307
| Rel.Std.Dev. |  | 0.7288% | 2.793% | 2.126% | 3.836% | 1.851% | 0.9293% | 1.033% | 1.757% | 2.678% | 0.1894% | 0.824% | 2.061%
| PDM3 |  | -1.288% | -0.03613% | 0.8081% | 2.031% | 1.827% | -2.714% | -4.577% | 1.681% | -14.25% | 1.208% | 9.454% | -11.62%
| Z-Score (Absolute) |  | 0.9 | 0.04 | 0.04 | 0.45 | 0.37 | 1.47 | 2.23 | 0.31 | 6.16 | 0.12 | 3.47 | 5.1
| Analytical results for Pr in OREAS 465 (Certified Value 3670 ppm) | [15 empty] | Certified
Pr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 3894.5207 | 3545.5 | 3456.75 | 2850 | 3875.19 | 3716.115 | 3411.5 | 3540 |  | 1
| 1 | 2 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 3949.703 | 3589.4 | 3474.19 | 2780 | 3763.26 | 3656.8204 | 3494 | 3520 |  | 0
| 1 | 3 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 3922.1119 | 3549.4 | 3423.85 | 2850 | 3928.57 | 3790.0336 | 3519.2 | 3550 |  | 16
| 1 | 4 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 3962.7181 | 3579.5 | 3708.22 | 2900 | 4070.74 | 3702.223 | 3499.2 | 3640 |  | 3669.9075
| 1 | 5 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 3998.6697 | 3623.5 | 3497.64 | 2930 | 3859.07 | 3670.9189 | 3591.7 | 3760 |  | 93
| 1 | 6 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 4014.7873 | 3608.6 | 3425.41 | 2820 | 3980.35 | 3695.2048 | 3556.2 | 3570
| Mean | [6 empty] | 3957.0851 | 3582.65 | 3497.6767 | 2855 | 3912.8633 | 3705.2193 | 3511.9667 | 3596.6667
| Median | [6 empty] | 3956.2105 | 3584.45 | 3465.47 | 2850 | 3901.88 | 3698.7139 | 3509.2 | 3560
| Std Dev. | [6 empty] | 45.3486 | 31.2515 | 106.9841 | 53.9444 | 106.2265 | 46.777 | 61.5559 | 90.037
| Rel.Std.Dev. | [6 empty] | 1.146% | 0.8723% | 3.059% | 1.889% | 2.715% | 1.262% | 1.753% | 2.503%
| PDM3 | [6 empty] | 7.825% | -2.378% | -4.693% | -22.21% | 6.62% | 0.9622% | -4.304% | -1.996%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 2.07 | 0.05 | 0.64 | 3.88 | 1.83 | 0.71 | 0.33 | 0.05
| Analytical results for Rb in OREAS 465 (Certified Value 0.43 ppm) | [15 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
2
| 1 | 1 | 0.4 | 0.4 | 0.4 | 0.5 | 0.4 | 0.53794 | 0.4 | 0.9 | 0.39 | 0.83 | 1.0567 | < 0.5 | 0.4 |  | 1
| 1 | 2 | 0.5 | 0.5 | 0.4 | 0.4 | 0.4 | 0.51145 | 0.4 | 0.8 | 0.38 | 0.36 | 1.0508 | < 0.5 | 0.4 |  | 19
| 1 | 3 | 0.2 | 0.4 | 0.4 | 0.4 | 0.4 | 0.50884 | 0.4 | 0.9 | 0.37 | 0.43 | 1.0015 | < 0.5 | 0.4 |  | 16
| 1 | 4 | 0.5 | 0.4 | 0.4 | 0.4 | 0.4 | 0.46025 | 0.5 | 0.7 | 0.35 | 0.52 | 0.99891 | < 0.5 | 0.4 |  | 0.42577
| 1 | 5 | 0.2 | 0.5 | 0.5 | 0.5 | 0.4 | 0.48918 | 0.4 | 1 | 0.43 | 0.82 | 0.99641 | < 0.5 | 0.4 |  | 94
| 1 | 6 | 0.3 | 0.4 | 0.4 | 0.5 | 0.4 | 0.55379 | 0.4 | 0.8 | 0.36 | 0.57 | 0.95786 | < 0.5 | 0.4
| Mean |  | 0.35 | 0.43333 | 0.41667 | 0.45 | 0.4 | 0.51024 | 0.41667 | 0.85 | 0.38 | 0.58833 | 1.0104 |  | 0.4
| Median |  | 0.35 | 0.4 | 0.4 | 0.45 | 0.4 | 0.51015 | 0.4 | 0.85 | 0.375 | 0.545 | 1.0002 |  | 0.4
| Std Dev. |  | 0.13784 | 0.05164 | 0.040825 | 0.054772 | 6.0809e-17 | 0.033496 | 0.040825 | 0.10488 | 0.028284 | 0.19712 | 0.03725 |  | 6.0809e-17
| Rel.Std.Dev. |  | 39.38% | 11.92% | 9.798% | 12.17% | 1.52e-14% | 6.565% | 9.798% | 12.34% | 7.443% | 33.5% | 3.687% |  | 1.52e-14%
| PDM3 |  | -17.8% | 1.777% | -2.138% | 5.691% | -6.052% | 19.84% | -2.138% | 99.64% | -10.75% | 38.18% | 137.3% |  | -6.052%
| Z-Score (Absolute) |  | 0.1 | 0.1 | 0.29 | 0.48 | 0.67 | 1.87 | 0.29 | 9.73 | 1.14 | 3.68 | 13.44 | NA | 0.67
| Analytical results for Re in OREAS 465 (Certified Value < 0.05 ppm) | [15 empty] | Certified
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 17 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*MS |  |  |  | 3
3
| 1 | 1 | 0.048 | 0.009 | 0.022 | 0.013 | 0.02 | < 0.002 | < 0.02 | 0.007 | 0.004 | < 0.05 | < 0.1 |  |  |  | 1
| 1 | 2 | 0.048 | 0.019 | 0.017 | 0.018 | 0.022 | < 0.002 | < 0.02 | 0.009 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 20
| 1 | 3 | 0.048 | 0.003 | 0.02 | 0.017 | 0.02 | < 0.002 | < 0.02 | 0.008 | 0.025 | < 0.05 | < 0.1 |  |  |  | 16
| 1 | 4 | 0.052 | 0.012 | 0.024 | 0.019 | 0.021 | < 0.002 | < 0.02 | 0.01 | 0.007 | < 0.05 | < 0.1 |  |  |  | < 0.05
| 1 | 5 | 0.049 | 0.01 | 0.023 | 0.021 | 0.019 | < 0.002 | < 0.02 | 0.011 | < 0.002 | < 0.05 | < 0.1 |  |  |  | 95
| 1 | 6 | 0.075 | 0.007 | 0.02 | 0.019 | 0.02 | < 0.002 | 0.02 | 0.01 | 0.004 | < 0.05 | < 0.1
| Mean |  | 0.053333 | 0.01 | 0.021 | 0.017833 | 0.020333 |  | 0.02 | 0.0091667 | 0.01
| Median |  | 0.0485 | 0.0095 | 0.021 | 0.0185 | 0.02 |  | 0.02 | 0.0095 | 0.0055
| Std Dev. |  | 01: 0.010727; 02: 0.0053666; 03: 0.0025298; 05: 0.0027142; 06: 0.0010328; 09: 0.001472; 10: 0.0101
| Rel.Std.Dev. |  | 01: 20.11%; 02: 53.67%; 03: 12.05%; 05: 15.22%; 06: 5.079%; 09: 16.06%; 10: 101%
PDM3
| Z-Score (Absolute) |  | 3.6 | 1.09 | 0.26 | 0.13 | 0.17 | NA | 0.13 | 1.2 | 1.71 | NA | NA
| Analytical results for S in OREAS 465 (Indicative Value 1103 ppm) | [15 empty] | Indicative
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 08 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 600 | 800 | 600 | 700 | 600 | 1700 | 1700 | 1750 | 1104 | 1200 | 1950 | 600 |  |  | 1
| 1 | 2 | 500 | 700 | 500 | 700 | 700 | 1700 | 1600 | 1770 | 1130 | 1200 | 1950 | 650 |  |  | 11
| 1 | 3 | 500 | 800 | 500 | 700 | 700 | 1700 | 1500 | 1790 |  | 1100 | 1900 | 550 |  |  | 16
| 1 | 4 | 600 | 800 | 600 | 700 | 700 | 1700 | 1600 | 1800 | 1262 | 1200 | 1950 | 700 |  |  | 1103.0972
| 1 | 5 | 600 | 800 | 500 | 800 | 700 | 1700 | 1700 | 1790 | 1065 | 1100 | 1950 | 650 |  |  | 29
| 1 | 6 | 500 | 800 | 600 | 800 | 600 | 1700 | 1800 | 1800 | 1083 | 1100 | 1900 | 700
| Mean |  | 550 | 783.3333 | 550 | 733.3333 | 666.6667 | 1700 | 1650 | 1783.3333 | 1128.8 | 1150 | 1933.3333 | 641.6667
| Median |  | 550 | 800 | 550 | 700 | 700 | 1700 | 1650 | 1790 | 1104 | 1150 | 1950 | 650
| Std Dev. |  | 54.7723 | 40.8248 | 54.7723 | 51.6398 | 51.6398 | 0 | 104.8809 | 19.6638 | 78.3052 | 54.7723 | 25.8199 | 58.4523
| Rel.Std.Dev. |  | 9.959% | 5.212% | 9.959% | 7.042% | 7.746% | 0% | 6.356% | 1.103% | 6.937% | 4.763% | 1.336% | 9.109%
| PDM3 |  | -50.14% | -28.99% | -50.14% | -33.52% | -39.56% | 54.11% | 49.58% | 61.67% | 2.33% | 4.252% | 75.26% | -41.83%
| Z-Score (Absolute) |  | 0.76 | 0.31 | 0.76 | 0.4 | 0.54 | 1.49 | 1.39 | 1.66 | 0.31 | 0.41 | 1.95 | 0.58
| Analytical results for Sb in OREAS 465 (Indicative Value 0.57 ppm) | [15 empty] | Indicative
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.25 | 0.25 | 0.22 | 0.17 | 0.27 | 1.1338 | 1.1 | 0.9 | 1.01 | 0.34 | < 0.2 |  |  |  | 1
| 1 | 2 | 0.32 | 0.12 | 0.19 | 0.21 | 0.24 | 1.0429 | 1.2 | 0.7 | 1.04 | 0.49 | < 0.2 |  |  |  | 21
| 1 | 3 | 0.26 | 0.27 | 0.19 | 0.2 | 0.24 | 1.1796 | 1.2 | < 0.5 | 1.07 | 0.47 | < 0.2 |  |  |  | 16
| 1 | 4 | 0.31 | 0.25 | 0.22 | 0.21 | 0.26 | 0.91385 | 1.2 | 0.9 | 1.01 | 0.42 | < 0.2 |  |  |  | 0.5746
| 1 | 5 | 0.3 | 0.29 | 0.19 | 0.32 | 0.22 | 0.93232 | 1 | 0.7 | 1.07 | 0.74 | < 0.2 |  |  |  | 30
| 1 | 6 | 0.38 | 0.24 | 0.17 | 0.31 | 0.21 | 1.0235 | 1 | 0.9 | 1.03 | 0.52 | < 0.2
| Mean |  | 0.30333 | 0.23667 | 0.19667 | 0.23667 | 0.24 | 1.0377 | 1.1167 | 0.82 | 1.0383 | 0.49667
| Median |  | 0.305 | 0.25 | 0.19 | 0.21 | 0.24 | 1.0332 | 1.15 | 0.9 | 1.035 | 0.48
| Std Dev. |  | 0.046762 | 0.059889 | 0.019664 | 0.062503 | 0.022804 | 0.10586 | 0.098319 | 0.10954 | 0.027142 | 0.13486
| Rel.Std.Dev. |  | 15.42% | 25.31% | 9.999% | 26.41% | 9.501% | 10.2% | 8.805% | 13.36% | 2.614% | 27.15%
| PDM3 |  | -47.21% | -58.81% | -65.77% | -58.81% | -58.23% | 80.59% | 94.34% | 42.71% | 80.71% | -13.56%
| Z-Score (Absolute) |  | 0.36 | 0.51 | 0.75 | 0.6 | 0.59 | 2.35 | 2.64 | 1.54 | 2.35 | 0.36 | NA
| Analytical results for Sc in OREAS 465 (Certified Value 149 ppm) | [15 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES |  |  | 0
0
| 1 | 1 | 136.5 | 146 | 149.5 | 154 | 144.5 | 144.7956 | 153 | 137 | 72 | 155.2816 | 149 | 165 |  |  | 1
| 1 | 2 | 145 | 152.5 | 147 | 153.5 | 146 | 145.9367 | 157 | 132 | 75.6 | 157.5553 | 148 | 163 |  |  | 15
| 1 | 3 | 112 | 147.5 | 146.5 | 161 | 138 | 146.5445 | 150 | 139 | 73.2 | 155.8924 | 139 | 163 |  |  | 16
| 1 | 4 | 141 | 145.5 | 156 | 160 | 138 | 147.7527 | 154 | 148 | 75 | 155.9534 | 132 | 164 |  |  | 148.5346
| 1 | 5 | 120 | 147.5 | 157 | 148.5 | 146 | 149.4149 | 149 | 153 | 73.5 | 157.6896 | 134 | 162 |  |  | 96
| 1 | 6 | 147 | 147.5 | 145.5 | 145 | 144 | 147.4697 | 156 | 139 | 74.2 | 157.7997 | 139 | 168
| Mean |  | 133.5833 | 147.75 | 150.25 | 153.6667 | 142.75 | 146.9857 | 153.1667 | 141.3333 | 73.9167 | 156.6953 | 140.1667 | 164.1667
| Median |  | 138.75 | 147.5 | 148.25 | 153.75 | 144.25 | 147.0071 | 153.5 | 139 | 73.85 | 156.7543 | 139 | 163.5
| Std Dev. |  | 14.3123 | 2.485 | 5.0274 | 6.2583 | 3.765 | 1.6009 | 3.1885 | 7.7115 | 1.2999 | 1.1083 | 7.0261 | 2.137
| Rel.Std.Dev. |  | 10.71% | 1.682% | 3.346% | 4.073% | 2.637% | 1.089% | 2.082% | 5.456% | 1.759% | 0.7073% | 5.013% | 1.302%
| PDM3 |  | -10.07% | -0.5283% | 1.155% | 3.455% | -3.894% | -1.043% | 3.118% | -4.848% | -50.24% | 5.494% | -5.634% | 10.52%
| Z-Score (Absolute) |  | 0.93 | 0.01 | 0.35 | 0.7 | 0.43 | 0.01 | 0.65 | 0.58 | 7.57 | 1.02 | 0.7 | 1.79
| Analytical results for Se in OREAS 465 (Indicative Value 35.4 ppm) | [15 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
1
| 1 | 1 | 14 | < 50 | 19 | 32 | 43 | 10.1 | 40 | 36 | < 0.5 | 4.9 | 35 |  |  |  | 1
| 1 | 2 | 14 | < 50 | 19 | 32 | 45 | 10.7655 | 40 | 43 | < 0.5 | 3.1 | 35 |  |  |  | 25
| 1 | 3 | 13 | < 50 | 18 | 33 | 42 | 9.2941 | 40 | 37 | < 0.5 | 4.9 | 35 |  |  |  | 16
| 1 | 4 | 14 | < 50 | 24 | 33 | 43 | 9.5819 | 39 | 41 | < 0.5 | 4 | 35 |  |  |  | 35.4167
| 1 | 5 | 14 | < 50 | 25 | 36 | 44 | 9.8006 | 39 | 40 | < 0.5 | < 0.5 | 35 |  |  |  | 31
| 1 | 6 | 26 | < 50 | 23 | 38 | 44 | 11.2686 | 39 | 38 | < 0.5 | < 0.5 | 35
| Mean |  | 15.8333 |  | 21.3333 | 34 | 43.5 | 10.1351 | 39.5 | 39.1667 |  | 4.225 | 35
| Median |  | 14 |  | 21 | 33 | 43.5 | 9.9503 | 39.5 | 39 |  | 4.45 | 35
| Std Dev. |  | 4.9967 |  | 3.0111 | 2.4495 | 1.0488 | 0.74948 | 0.54772 | 2.6394 |  | 0.86168 | 0
| Rel.Std.Dev. |  | 31.56% |  | 14.11% | 7.204% | 2.411% | 7.395% | 1.387% | 6.739% |  | 20.39% | 0%
| PDM3 |  | -55.29% |  | -39.76% | -4% | 22.82% | -71.38% | 11.53% | 10.59% |  | -88.07% | -1.176%
| Z-Score (Absolute) |  | 1.43 | NA | 0.9 | 0 | 0.67 | 1.69 | 0.39 | 0.37 | NA | 2.11 | 0.07
| Analytical results for Sm in OREAS 465 (Certified Value 1307 ppm) | [15 empty] | Certified
Sm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1387.603 | 1300 | 1241.4 | 557 | 1329.69 | 1221.6651 | 1201 | 1350 |  | 1
| 1 | 2 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1418.9781 | 1332 | 1230.7 | 551 | 1323.98 | 1178.5386 | 1245.8 | 1340 |  | 0
| 1 | 3 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1465.9338 | 1324.4 | 1219.9 | 581 | 1391.98 | 1227.4241 | 1254.4 | 1350 |  | 16
| 1 | 4 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1417.9958 | 1291.9 | 1312.7 | 597 | 1398.13 | 1171.3579 | 1253.2 | 1370 |  | 1306.679
| 1 | 5 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1452.6193 | 1325.8 | 1261.3 | 551 | 1373.78 | 1183.0667 | 1268.9 | 1410 |  | 97
| 1 | 6 | > 1000 | > 1000 | > 1000 | > 1000 | > 1000 | 1432.2242 | 1318.1 | 1228.6 | 567 | 1339.88 | 1205.7929 | 1266.1 | 1340
| Mean | [6 empty] | 1429.2257 | 1315.3667 | 1249.1 | 567.3333 | 1359.5733 | 1197.9742 | 1248.2333 | 1360
| Median | [6 empty] | 1425.6012 | 1321.25 | 1236.05 | 562 | 1356.83 | 1194.4298 | 1253.8 | 1350
| Std Dev. | [6 empty] | 27.8112 | 15.8834 | 34.2477 | 18.4788 | 32.5157 | 23.6563 | 24.6817 | 26.8328
| Rel.Std.Dev. | [6 empty] | 1.946% | 1.208% | 2.742% | 3.257% | 2.392% | 1.975% | 1.977% | 1.973%
| PDM3 | [6 empty] | 9.378% | 0.6649% | -4.407% | -56.58% | 4.048% | -8.319% | -4.473% | 4.081%
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 1.27 | 0.29 | 0.39 | 6.18 | 0.67 | 0.72 | 0.29 | 0.68
| Analytical results for Sn in OREAS 465 (Indicative Value 48.8 ppm) | [15 empty] | Indicative
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 1
1
| 1 | 1 | 18.1 | 26.9 | 22.4 | 18.7 | 23.5 | 92.1303 | 93.9 | 84 | 75.1 | 61.9 | 2 |  |  |  | 1
| 1 | 2 | 21.4 | 22.4 | 18.6 | 17.2 | 22.7 | 93.247 | 96.5 | 78 | 75.7 | 72.9 | 3 |  |  |  | 10
| 1 | 3 | 16.3 | 28 | 19.5 | 20.7 | 21.3 | 94.9483 | 95.5 | 75 | 74.6 | 62.7 | 2 |  |  |  | 16
| 1 | 4 | 24 | 28.3 | 22.1 | 19.3 | 23.8 | 92.9094 | 96.7 | 85 | 76.1 | 62.3 | 3 |  |  |  | 48.7534
| 1 | 5 | 18.9 | 27.9 | 20.2 | 36.2 | 22.4 | 93.1175 | 93 | 94 | 74.4 | 88.9 | 5 |  |  |  | 32
| 1 | 6 | 21.8 | 27.4 | 19 | 34.6 | 20.1 | 94.773 | 94.6 | 89 | 75.7 | 73.1 | 4
| Mean |  | 20.0833 | 26.8167 | 20.3 | 24.45 | 22.3 | 93.5209 | 95.0333 | 84.1667 | 75.2667 | 70.3 | 3.1667
| Median |  | 20.15 | 27.65 | 19.85 | 20 | 22.55 | 93.1822 | 95.05 | 84.5 | 75.4 | 67.8 | 3
| Std Dev. |  | 2.8167 | 2.2194 | 1.605 | 8.5708 | 1.3928 | 1.1091 | 1.4665 | 6.969 | 0.67725 | 10.5155 | 1.169
| Rel.Std.Dev. |  | 14.02% | 8.276% | 7.906% | 35.05% | 6.246% | 1.186% | 1.543% | 8.28% | 0.8998% | 14.96% | 36.92%
| PDM3 |  | -58.81% | -45% | -58.36% | -49.85% | -54.26% | 91.82% | 94.93% | 72.64% | 54.38% | 44.2% | -93.5%
| Z-Score (Absolute) |  | 0.21 | 0 | 0.2 | 0.09 | 0.15 | 1.81 | 1.85 | 1.55 | 1.31 | 1.17 | 0.67
| Analytical results for Sr in OREAS 465 (Certified Value 0.505 wt.%) | [15 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 3
3
| 1 | 1 | 0.546 | 0.501 | 0.288 | 0.321 | 0.264 | 0.5085 | 0.49794 | 0.54666 | 0.471 | 0.51811 | 0.50651 | 0.4691 | 0.5135 | 0.456 | 1
| 1 | 2 | 0.559 | 0.467 | 0.285 | 0.309 | 0.268 | 0.51865 | 0.51326 | 0.5408 | 0.473 | 0.49378 | 0.505 | 0.4774 | 0.5175 | 0.464 | 1
| 1 | 3 | 0.541 | 0.5 | 0.282 | 0.322 | 0.275 | 0.52503 | 0.49733 | 0.55019 | 0.477 | 0.49953 | 0.50383 | 0.4848 | 0.513 | 0.465 | 16
| 1 | 4 | 0.556 | 0.505 | 0.292 | 0.312 | 0.271 | 0.51129 | 0.49909 | 0.55955 | 0.486 | 0.50622 | 0.5013 | 0.4672 | 0.513 | 0.466 | 0.50533
| 1 | 5 | 0.542 | 0.501 | 0.289 | 0.516 | 0.274 | 0.52546 | 0.49921 | 0.55267 | 0.487 | 0.50745 | 0.50084 | 0.4754 | 0.515 | 0.498 | 98
| 1 | 6 | 0.521 | 0.496 | 0.293 | 0.531 | 0.261 | 0.52241 | 0.50818 | 0.55486 | 0.487 | 0.50691 | 0.5012 | 0.4801 | 0.5115 | 0.468
| Mean |  | 0.54417 | 0.495 | 0.28817 | 0.38517 | 0.26883 | 0.51856 | 0.5025 | 0.55079 | 0.48017 | 0.50533 | 0.50311 | 0.47567 | 0.51392 | 0.4695
| Median |  | 0.544 | 0.5005 | 0.2885 | 0.3215 | 0.2695 | 0.52053 | 0.49915 | 0.55143 | 0.4815 | 0.50656 | 0.50257 | 0.4764 | 0.51325 | 0.4655
| Std Dev. |  | 0.013527 | 0.014014 | 0.0041673 | 0.10737 | 0.0055648 | 0.0071864 | 0.0066033 | 0.006543 | 0.0073869 | 0.0082227 | 0.0023542 | 0.0066476 | 0.0020837 | 0.014557
| Rel.Std.Dev. |  | 2.486% | 2.831% | 1.446% | 27.88% | 2.07% | 1.386% | 1.314% | 1.188% | 1.538% | 1.627% | 0.4679% | 1.398% | 0.4054% | 3.1%
| PDM3 |  | 7.686% | -2.044% | -42.97% | -23.78% | -46.8% | 2.618% | -0.5593% | 8.996% | -4.979% | 0.0007261% | -0.4381% | -5.87% | 1.7% | -7.09%
| Z-Score (Absolute) |  | 1.22 | 0.03 | 6.09 | 3.32 | 6.64 | 0.49 | 0.03 | 1.4 | 0.61 | 0.11 | 0.04 | 0.74 | 0.35 | 1.08
| Analytical results for Ta in OREAS 465 (Indicative Value 50 ppm) | [15 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  |  | 0
0
| 1 | 1 | 4.35 | 5.03 | 3.48 | 5.26 | 59.8519 | 55.2 | 13.2 | 55.1 | 22.96 | 0.4 |  |  |  |  | 1
| 1 | 2 | 5.85 | 3.78 | 2.6 | 5.16 | 53.4843 | 55.5 | 4.1 | 54.7 | 39.46 | 0.6 |  |  |  |  | 27
| 1 | 3 | 3.54 | 4.19 | 3.99 | 5.35 | 56.028 | 52.9 | 1.4 | 54.8 | 36.64 | 0.4 |  |  |  |  | 16
| 1 | 4 | 7.12 | 4.39 | 3.41 | 6.33 | 57.2286 | 57 | 19.7 | 56.7 | 24.75 | 0.3 |  |  |  |  | 50.1017
| 1 | 5 | 4.2 | 3.72 | 9.88 | 6.18 | 51.1724 | 54.6 | 7.3 | 55.5 | 47.12 | 0.6 |  |  |  |  | 33
| 1 | 6 | 5.23 | 3.2 | 9.15 | 4.73 | 51.7347 | 56.3 | 9.7 | 54.4 | 41.11 | 0.4
| Mean |  | 5.0483 | 4.0517 | 5.4183 | 5.5017 | 54.9166 | 55.25 | 9.2333 | 55.2 | 35.34 | 0.45
| Median |  | 4.79 | 3.985 | 3.735 | 5.305 | 54.7561 | 55.35 | 8.5 | 54.95 | 38.05 | 0.4
| Std Dev. |  | 1.3 | 0.63244 | 3.2126 | 0.62288 | 3.3803 | 1.4265 | 6.5872 | 0.82462 | 9.5513 | 0.12247
| Rel.Std.Dev. |  | 25.75% | 15.61% | 59.29% | 11.32% | 6.155% | 2.582% | 71.34% | 1.494% | 27.03% | 27.22%
| PDM3 |  | -89.92% | -91.91% | -89.19% | -89.02% | 9.61% | 10.28% | -81.57% | 10.18% | -29.46% | -99.1%
| Z-Score (Absolute) |  | 0.31 | 0.44 | 0.26 | 0.25 | 6.27 | 6.31 | 0.25 | 6.26 | 3.69 | 0.91
| Analytical results for Tb in OREAS 465 (Certified Value 57 ppm) | [15 empty] | Certified
Tb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
1
| 1 | 1 | 58 | 54.8 | 53.1 | 54.6 | 62.5 | 58.1635 | 61.6 | 57.24 | 47.3 | 59.79 | 86.9937 | 53.85 | 57.2 |  | 1
| 1 | 2 | 58 | 50.6 | 49.3 | 53.8 | 63.4 | 59.0905 | 61.7 | 56.59 | 48.8 | 58.17 | 85.6789 | 54.1 | 55 |  | 3
| 1 | 3 | 52.8 | 56 | 50.9 | 56.1 | 60.3 | 59.6634 | 58.9 | 56.7 | 48 | 62.61 | 86.7447 | 54.15 | 53.7 |  | 16
| 1 | 4 | 58.5 | 56.2 | 53.3 | 56.2 | 60.9 | 57.5953 | 62.2 | 58.58 | 44.3 | 63.21 | 85.3252 | 53.9 | 56.9 |  | 57.0775
| 1 | 5 | 53.4 | 58 | 53 | 56.4 | 60.2 | 57.9521 | 60.9 | 56.58 | 48.9 | 60.33 | 85.2972 | 55.2 | 59.4 |  | 99
| 1 | 6 | 57.6 | 55 | 52.3 | 56.7 | 57.7 | 58.6832 | 64.2 | 56.21 | 48 | 60.32 | 84.9943 | 54.35 | 54.7
| Mean |  | 56.3833 | 55.1 | 51.9833 | 55.6333 | 60.8333 | 58.5247 | 61.5833 | 56.9833 | 47.55 | 60.7383 | 85.839 | 54.2583 | 56.15
| Median |  | 57.8 | 55.5 | 52.65 | 56.15 | 60.6 | 58.4233 | 61.65 | 56.645 | 48 | 60.325 | 85.5021 | 54.125 | 55.95
| Std Dev. |  | 2.5663 | 2.4811 | 1.5804 | 1.157 | 1.9937 | 0.76993 | 1.7268 | 0.84982 | 1.6979 | 1.8681 | 0.83071 | 0.4954 | 2.0811
| Rel.Std.Dev. |  | 4.551% | 4.503% | 3.04% | 2.08% | 3.277% | 1.316% | 2.804% | 1.491% | 3.571% | 3.076% | 0.9678% | 0.913% | 3.706%
| PDM3 |  | -1.216% | -3.465% | -8.925% | -2.53% | 6.58% | 2.536% | 7.894% | -0.1649% | -16.69% | 6.414% | 50.39% | -4.939% | -1.625%
| Z-Score (Absolute) |  | 0 | 0.4 | 1.39 | 0.24 | 1.4 | 0.67 | 1.64 | 0.09 | 2.58 | 1.37 | 9.28 | 0.67 | 0.07
| Analytical results for Te in OREAS 465 (Certified Value < 1 ppm) | [15 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 0.58 | 0.32 | 0.63 | 0.36 | 0.38 | 0.47536 | < 1 | < 0.5 | 0.77 | 0.6 | 0.2 |  |  |  | 1
| 1 | 2 | 0.59 | 0.37 | 0.58 | 0.36 | 0.36 | 0.41185 | < 1 | < 0.5 | 0.77 | 0.5 | 0.4 |  |  |  | 17
| 1 | 3 | 0.5 | 0.3 | 0.62 | 0.41 | 0.39 | 0.45528 | < 1 | < 0.5 | 0.74 | 0.5 | 0.2 |  |  |  | 16
| 1 | 4 | 0.62 | 0.35 | 0.38 | 0.39 | 0.42 | 0.48238 | < 1 | < 0.5 | 0.78 | 0.5 | 0.2 |  |  |  | < 1
| 1 | 5 | 0.56 | 0.34 | 0.41 | 0.44 | 0.36 | 0.47409 | < 1 | < 0.5 | 0.78 | 0.6 | 0.2 |  |  |  | 100
| 1 | 6 | 0.64 | 0.32 | 0.32 | 0.41 | 0.34 | 0.44 | < 1 | < 0.5 | 0.7 | 0.5 | 0.4
| Mean |  | 0.58167 | 0.33333 | 0.49 | 0.395 | 0.375 | 0.45649 |  |  | 0.75667 | 0.53333 | 0.26667
| Median |  | 0.585 | 0.33 | 0.495 | 0.4 | 0.37 | 0.46468 |  |  | 0.77 | 0.5 | 0.2
| Std Dev. |  | 0.04916 | 0.025033 | 0.13565 | 0.031464 | 0.028107 | 0.026839 |  |  | 0.031411 | 0.05164 | 0.10328
| Rel.Std.Dev. |  | 8.452% | 7.51% | 27.68% | 7.966% | 7.495% | 5.879% |  |  | 4.151% | 9.682% | 38.73%
PDM3
| Z-Score (Absolute) |  | 1.04 | 1.02 | 0.28 | 0.51 | 0.67 | 0 | NA | NA | 2.58 | 0.64 | 1.57
| Analytical results for Th in OREAS 465 (Certified Value 805 ppm) | [15 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS |  |  |  | 0
0
| 1 | 1 | 1130 | 680 | 650 | 680 | 580 | 836.5 | 857.8 | 118 | 894.34 | 849.793 | 834 |  |  |  | 1
| 1 | 2 | 1240 | 690 | 640 | 650 | 600 | 833.8146 | 835.3 | 117 | 883.46 | 851.2585 | 819 |  |  |  | 0
| 1 | 3 | 1010 | 650 | 640 | 690 | 600 | 856.1516 | 820.3 | 123 | 905.7 | 850.9998 | 825 |  |  |  | 16
| 1 | 4 | 1210 | 680 | 650 | 660 | 590 | 832.6305 | 890.6 | 122 | 941.81 | 853.8043 | 853 |  |  |  | 804.6092
| 1 | 5 | 1050 | 680 | 650 | 880 | 600 | 836.8645 | 857.4 | 120 | 908.16 | 851.9679 | 868 |  |  |  | 101
| 1 | 6 | 186.5 | 690 | 650 | 5.5 | 570 | 851.5361 | 850.4 | 123 | 911.29 | 852.6716 | 833
| Mean |  | 971.0833 | 678.3333 | 646.6667 | 594.25 | 590 | 841.2496 | 851.9667 | 120.5 | 907.46 | 851.7492 | 838.6667
| Median |  | 1090 | 680 | 650 | 670 | 595 | 836.6823 | 853.9 | 121 | 906.93 | 851.6132 | 833.5
| Std Dev. |  | 394.4528 | 14.7196 | 5.164 | 300.743 | 12.6491 | 9.9924 | 23.8289 | 2.5884 | 19.7297 | 1.3965 | 18.4029
| Rel.Std.Dev. |  | 40.62% | 2.17% | 0.7986% | 50.61% | 2.144% | 1.188% | 2.797% | 2.148% | 2.174% | 0.164% | 2.194%
| PDM3 |  | 20.69% | -15.69% | -19.63% | -26.14% | -26.67% | 4.554% | 5.886% | -85.02% | 12.78% | 5.859% | 4.233%
| Z-Score (Absolute) |  | 1.54 | 0.85 | 1.02 | 0.67 | 1.32 | 0.01 | 0.07 | 3.82 | 0.37 | 0.07 | 0
| Analytical results for Ti in OREAS 465 (Indicative Value 1.3 wt.%) | [15 empty] | Indicative
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 22 |  |  | wt.%
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 0.813 | 0.98 | 0.465 | 0.436 | 0.439 | 2.7674 | 1.98 | 1.827 | 1.93 | 2.1761 | 0.65 | 1.29 |  |  | 1
| 1 | 2 | 0.894 | 0.99 | 0.475 | 0.405 | 0.421 | 2.7247 | 2.06 | 1.6434 | 1.98 | 1.5632 | 0.67 | 1.41 |  |  | 4
| 1 | 3 | 0.726 | 1 | 0.456 | 0.454 | 0.482 | 2.5325 | 2.04 | 1.4899 | 1.97 | 1.5827 | 0.63 | 1.21 |  |  | 16
| 1 | 4 | 0.885 | 1 | 0.455 | 0.425 | 0.46 | 2.5885 | 2.05 | 2.1091 | 1.97 | 2.2163 | 0.67 | 1.31 |  |  | 1.3009
| 1 | 5 | 0.769 | 1.02 | 0.464 | 0.949 | 0.436 | 2.4922 | 1.96 | 2.2147 | 1.94 | 1.7016 | 0.66 | 1.34 |  |  | 29
| 1 | 6 | 0.823 | 0.97 | 0.454 | 0.958 | 0.42 | 2.5359 | 1.98 | 2.1478 | 1.93 | 1.7791 | 0.66 | 1.36
| Mean |  | 0.81833 | 0.99333 | 0.4615 | 0.6045 | 0.443 | 2.6069 | 2.0117 | 1.9053 | 1.9533 | 1.8365 | 0.65667 | 1.32
| Median |  | 0.818 | 0.995 | 0.46 | 0.445 | 0.4375 | 2.5622 | 2.01 | 1.9681 | 1.955 | 1.7404 | 0.66 | 1.325
| Std Dev. |  | 0.065084 | 0.017512 | 0.0081179 | 0.27082 | 0.024033 | 0.1129 | 0.043089 | 0.29777 | 0.022509 | 0.28989 | 0.015055 | 0.068118
| Rel.Std.Dev. |  | 7.953% | 1.763% | 1.759% | 44.8% | 5.425% | 4.331% | 2.142% | 15.63% | 1.152% | 15.78% | 2.293% | 5.16%
| PDM3 |  | -37.1% | -23.64% | -64.53% | -53.53% | -65.95% | 100.4% | 54.63% | 46.46% | 50.15% | 41.17% | -49.52% | 1.467%
| Z-Score (Absolute) |  | 0.33 | 0.16 | 0.68 | 0.54 | 0.7 | 1.42 | 0.84 | 0.73 | 0.78 | 0.67 | 0.49 | 0.16
| Analytical results for Tl in OREAS 465 (Certified Value 0.087 ppm) | [15 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 3
3
| 1 | 1 | 0.09 | < 0.02 | 0.1 | 0.12 | 0.09 | 0.070861 | 0.07 | < 0.2 | 0.08 | 0.05 | 0.24607 | 0.1 | < 0.1 |  | 1
| 1 | 2 | 0.09 | 0.04 | 0.08 | 0.09 | 0.09 | 0.078386 | 0.07 | < 0.2 | 0.08 | 0.05 | 0.25774 | 0.1 | < 0.1 |  | 18
| 1 | 3 | 0.09 | < 0.02 | 0.08 | 0.09 | 0.09 | 0.079224 | 0.09 | < 0.2 | 0.09 | 0.07 | 0.2885 | 0.1 | < 0.1 |  | 16
| 1 | 4 | 0.09 | < 0.02 | 0.1 | 0.09 | 0.08 | 0.08673 | 0.1 | < 0.2 | 0.09 | 0.06 | 0.2991 | 0.05 | < 0.1 |  | 0.08732
| 1 | 5 | 0.09 | < 0.02 | 0.09 | 0.08 | 0.09 | 0.090205 | 0.08 | < 0.2 | 0.09 | 0.05 | 0.23016 | 0.1 | < 0.1 |  | 102
| 1 | 6 | 0.12 | < 0.02 | 0.09 | 0.09 | 0.08 | 0.067976 | 0.08 | < 0.2 | 0.07 | 0.04 | 0.21213 | 0.05 | < 0.1
| Mean |  | 0.095 | 0.04 | 0.09 | 0.093333 | 0.086667 | 0.078897 | 0.081667 |  | 0.083333 | 0.053333 | 0.25562 | 0.083333
| Median |  | 0.09 | 0.04 | 0.09 | 0.09 | 0.09 | 0.078805 | 0.08 |  | 0.085 | 0.05 | 0.25191 | 0.1
| Std Dev. |  | 0.012247 |  | 0.0089443 | 0.013663 | 0.005164 | 0.0086421 | 0.01169 |  | 0.008165 | 0.010328 | 0.033489 | 0.02582
| Rel.Std.Dev. |  | 12.89% |  | 9.938% | 14.64% | 5.958% | 10.95% | 14.31% |  | 9.798% | 19.36% | 13.1% | 30.98%
| PDM3 |  | 8.795% | -54.19% | 3.069% | 6.886% | -0.7487% | -9.647% | -6.475% |  | -4.566% | -38.92% | 192.7% | -4.566%
| Z-Score (Absolute) |  | 0.45 | 6.29 | 0.45 | 0.18 | 0 | 1.05 | 0.67 | NA | 0.45 | 4.5 | 22.79 | 1.8 | NA
| Analytical results for Tm in OREAS 465 (Certified Value 3.82 ppm) | [15 empty] | Certified
Tm | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 2
3
| 1 | 1 | 4.35 | 3.89 | 3.75 | 3.72 | 3.98 | 3.4145 | 3.6431 | 4.3 | 2.82 | 3.68 | 4.6353 | 3.9 | 3.48 |  | 1
| 1 | 2 | 4.41 | 3.32 | 3.55 | 3.67 | 4.05 | 3.5465 | 3.6416 | 4.2 | 2.71 | 3.66 | 4.7668 | 3.9 | 3.42 |  | 5
| 1 | 3 | 4.06 | 3.89 | 3.63 | 3.86 | 3.94 | 3.4884 | 3.7046 | 4.2 | 2.83 | 3.64 | 4.6222 | 3.85 | 3.4 |  | 16
| 1 | 4 | 4.42 | 3.92 | 4.05 | 3.84 | 3.93 | 3.4027 | 3.6524 | 4.4 | 2.94 | 3.63 | 4.3618 | 3.9 | 3.42 |  | 3.8238
| 1 | 5 | 4.05 | 3.89 | 4.09 | 3.71 | 4.16 | 3.542 | 3.7173 | 4.2 | 2.79 | 3.85 | 4.5437 | 3.9 | 3.58 |  | 103
| 1 | 6 | 3.79 | 3.93 | 3.87 | 3.7 | 4.03 | 3.5257 | 3.7238 | 4.2 | 2.82 | 3.56 | 4.1463 | 3.9 | 3.38
| Mean |  | 4.18 | 3.8067 | 3.8233 | 3.75 | 4.015 | 3.4866 | 3.6804 | 4.25 | 2.8183 | 3.67 | 4.5127 | 3.8917 | 3.4467
| Median |  | 4.205 | 3.89 | 3.81 | 3.715 | 4.005 | 3.5071 | 3.6785 | 4.2 | 2.82 | 3.65 | 4.583 | 3.9 | 3.42
| Std Dev. |  | 0.25409 | 0.23905 | 0.22006 | 0.079498 | 0.085499 | 0.063918 | 0.038771 | 0.083666 | 0.07414 | 0.09716 | 0.2234 | 0.020412 | 0.073394
| Rel.Std.Dev. |  | 6.079% | 6.28% | 5.756% | 2.12% | 2.129% | 1.833% | 1.053% | 1.969% | 2.631% | 2.647% | 4.95% | 0.5245% | 2.129%
| PDM3 |  | 9.315% | -0.4479% | -0.01206% | -1.93% | 5% | -8.818% | -3.749% | 11.15% | -26.29% | -4.022% | 18.02% | 1.775% | -9.863%
| Z-Score (Absolute) |  | 1.25 | 0.28 | 0 | 0.26 | 0.67 | 1.18 | 0.5 | 1.5 | 3.54 | 0.67 | 2.43 | 0.24 | 1.33
| Analytical results for U in OREAS 465 (Certified Value 12.6 ppm) | [15 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 1
2
| 1 | 1 | 12.1 | 12 | 12.4 | 12.6 | 14.4 | 12.5154 | 12.69 | 12.7 | 12.6 | 14.27 | 15.5683 | 12 | 11 |  | 1
| 1 | 2 | 12.4 | 11.6 | 12.2 | 12.4 | 14.2 | 12.3375 | 12.65 | 12.7 | 12.7 | 12.87 | 14.9518 | 11.9 | 10.6 |  | 6
| 1 | 3 | 11.8 | 12.3 | 12.3 | 13.2 | 13.7 | 12.0757 | 12.49 | 12.6 | 12.9 | 13.46 | 15.5599 | 11.8 | 10.8 |  | 16
| 1 | 4 | 12.5 | 12.1 | 13.7 | 13.1 | 13.7 | 12.0711 | 12.64 | 13.2 | 13 | 14.1 | 14.6331 | 11.8 | 11.1 |  | 12.6051
| 1 | 5 | 11.9 | 12.2 | 13.5 | 11.3 | 12.3 | 12.4132 | 12.81 | 13 | 12.8 | 15.02 | 14.7101 | 12 | 11.2 |  | 104
| 1 | 6 | 12.5 | 12 | 12.9 | 11.5 | 12 | 12.278 | 12.68 | 12.7 | 12.5 | 13.09 | 14.4413 | 11.7 | 10.8
| Mean |  | 12.2 | 12.0333 | 12.8333 | 12.35 | 13.3833 | 12.2818 | 12.66 | 12.8167 | 12.75 | 13.8017 | 14.9774 | 11.8667 | 10.9167
| Median |  | 12.25 | 12.05 | 12.65 | 12.5 | 13.7 | 12.3077 | 12.665 | 12.7 | 12.75 | 13.78 | 14.831 | 11.85 | 10.9
| Std Dev. |  | 0.30984 | 0.24221 | 0.64395 | 0.79687 | 0.99883 | 0.17991 | 0.10315 | 0.23166 | 0.18708 | 0.81042 | 0.48303 | 0.12111 | 0.22286
| Rel.Std.Dev. |  | 2.54% | 2.013% | 5.018% | 6.452% | 7.463% | 1.465% | 0.8148% | 1.807% | 1.467% | 5.872% | 3.225% | 1.021% | 2.041%
| PDM3 |  | -3.214% | -4.536% | 1.81% | -2.024% | 6.174% | -2.565% | 0.4353% | 1.678% | 1.149% | 9.492% | 18.82% | -5.858% | -13.4%
| Z-Score (Absolute) |  | 0.67 | 0.92 | 0.25 | 0.45 | 1.06 | 0.55 | 0 | 0.12 | 0.13 | 1.32 | 3.4 | 1.16 | 2.56
| Analytical results for V in OREAS 465 (Certified Value 427 ppm) | [15 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES |  | 0
0
| 1 | 1 | 411 | 360 | 225 | 256 | 216 | 470.287 | 452 | 481 | 434 | 521 | 423.554 | 341 | 475 |  | 1
| 1 | 2 | 406 | 386 | 227 | 248 | 218 | 477.5638 | 460 | 470 | 450 |  | 425.484 | 342 | 460 |  | 7
| 1 | 3 | 399 | 359 | 221 | 257 | 224 | 482.2051 | 454 | 487 | 450 | 520 | 423.838 | 340 | 475 |  | 16
| 1 | 4 | 398 | 369 | 234 | 246 | 220 | 469.8223 | 456 | 498 | 461 | 527 | 424.62 | 340 | 465 |  | 426.604
| 1 | 5 | 396 | 370 | 230 | 385 | 220 | 477.3051 | 451 | 500 | 455 | 514 | 425.926 | 325 | 460 |  | 105
| 1 | 6 | 402 | 365 | 236 | 404 | 214 | 474.4841 | 460 | 492 | 458 | 536 | 424.177 | 326 | 480
| Mean |  | 402 | 368.1667 | 228.8333 | 299.3333 | 218.6667 | 475.2779 | 455.5 | 488 | 451.3333 | 523.6 | 424.5998 | 335.6667 | 469.1667
| Median |  | 400.5 | 367 | 228.5 | 256.5 | 219 | 475.8946 | 455 | 489.5 | 452.5 | 521 | 424.3985 | 340 | 470
| Std Dev. |  | 5.6214 | 9.8268 | 5.6362 | 74.0855 | 3.5024 | 4.746 | 3.8859 | 11.2606 | 9.5429 | 8.3247 | 0.93742 | 7.9162 | 8.612
| Rel.Std.Dev. |  | 1.398% | 2.669% | 2.463% | 24.75% | 1.602% | 0.9986% | 0.8531% | 2.307% | 2.114% | 1.59% | 0.2208% | 2.358% | 1.836%
| PDM3 |  | -5.767% | -13.7% | -46.36% | -29.83% | -48.74% | 11.41% | 6.773% | 14.39% | 5.797% | 22.74% | -0.4698% | -21.32% | 9.977%
| Z-Score (Absolute) |  | 0.27 | 0.67 | 2.34 | 1.5 | 2.46 | 0.61 | 0.37 | 0.76 | 0.32 | 1.18 | 0 | 1.06 | 0.53
| Analytical results for W in OREAS 465 (Indicative Value 2.4 ppm) | [15 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 22 |  |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS |  |  |  | 2
2
| 1 | 1 | 2.1 | 1.8 | 2.4 | 1.8 | 2 | 3.6848 | 3.0233 | 2 | 3 | 2.6 | 1 |  |  |  | 1
| 1 | 2 | 2.2 | 1.7 | 2.3 | 1.7 | 2 | 3.4808 | 3.09 | 2 | 3.3 | 2.6 | 1.5 |  |  |  | 8
| 1 | 3 | 1.9 | 1.8 | 2.3 | 1.9 | 2.1 | 3.8348 | 3.1053 | 2 | 3.3 | 2.9 | 1 |  |  |  | 16
| 1 | 4 | 2.2 | 1.8 | 2.5 | 1.9 | 2 | 3.8489 | 3.2623 | 3 | 3 | 2.6 | 1.5 |  |  |  | 2.397
| 1 | 5 | 1.9 | 1.8 | 2.5 | 2.2 | 1.9 | 3.5055 | 3.2464 | 3 | 3.1 | 2.8 | 1.5 |  |  |  | 30
| 1 | 6 | 2.4 | 1.8 | 2.3 | 2.3 | 2 | 3.3948 | 3.2084 | 2 | 3.3 | 2.5 | 1.5
| Mean |  | 2.1167 | 1.7833 | 2.3833 | 1.9667 | 2 | 3.6249 | 3.1559 | 2.3333 | 3.1667 | 2.6667 | 1.3333
| Median |  | 2.15 | 1.8 | 2.35 | 1.9 | 2 | 3.5952 | 3.1569 | 2 | 3.2 | 2.6 | 1.5
| Std Dev. |  | 0.19408 | 0.040825 | 0.098319 | 0.23381 | 0.063246 | 0.19275 | 0.096689 | 0.5164 | 0.15055 | 0.15055 | 0.2582
| Rel.Std.Dev. |  | 9.169% | 2.289% | 4.125% | 11.89% | 3.162% | 5.317% | 3.064% | 22.13% | 4.754% | 5.646% | 19.36%
| PDM3 |  | -11.69% | -25.6% | -0.5684% | -17.95% | -16.56% | 51.23% | 31.66% | -2.654% | 32.11% | 11.25% | -44.37%
| Z-Score (Absolute) |  | 0.4 | 1.01 | 0.09 | 0.67 | 0.61 | 2.38 | 1.51 | 0 | 1.53 | 0.61 | 1.84
| Analytical results for Y in OREAS 465 (Certified Value 478 ppm) | [15 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  | 0
0
| 1 | 1 | 470 | 403 | > 500 | 260 | > 500 | 465.9175 | 484.6908 | 517.5 | 325 | 524.6 | 570.4125 | 466.9 | 480 |  | 1
| 1 | 2 | 480 | 450 | > 500 | 257 | > 500 | 474.5762 | 490.2528 | 515.8 | 328 | 489.47 | 568.0929 | 467.4 | 477 |  | 0
| 1 | 3 | 439 | 412 | > 500 | 266 | > 500 | 467.3534 | 492.4365 | 518.7 | 334 | 508.52 | 571.7413 | 464.1 | 491 |  | 16
| 1 | 4 | 480 | 412 | > 500 | 261 | > 500 | 467.487 | 487.5768 | 528 | 334 | 524.3 | 572.152 | 461.7 | 488 |  | 477.799
| 1 | 5 | 448 | 412 | > 500 | 419 | 473 | 471.5688 | 490.2522 | 520.1 | 316 | 538.75 | 572.4054 | 467.6 | 516 |  | 106
| 1 | 6 | 460 | 401 | 494 | 428 | 467 | 477.2609 | 493.3673 | 517.9 | 340 | 493.16 | 569.6351 | 465.3 | 486
| Mean |  | 462.8333 | 415 | 494 | 315.1667 | 470 | 470.694 | 489.7627 | 519.6667 | 329.5 | 513.1333 | 570.7399 | 465.5 | 489.6667
| Median |  | 465 | 412 | 494 | 263.5 | 470 | 469.5279 | 490.2525 | 518.3 | 331 | 516.41 | 571.0769 | 466.1 | 487
| Std Dev. |  | 16.9519 | 17.8438 |  | 84.0129 | 4.2426 | 4.5436 | 3.1976 | 4.3205 | 8.4321 | 19.4554 | 1.6796 | 2.2987 | 13.8948
| Rel.Std.Dev. |  | 3.663% | 4.3% |  | 26.66% | 0.9027% | 0.9653% | 0.6529% | 0.8314% | 2.559% | 3.791% | 0.2943% | 0.4938% | 2.838%
| PDM3 |  | -3.132% | -13.14% | 3.391% | -34.04% | -1.632% | -1.487% | 2.504% | 8.763% | -31.04% | 7.395% | 19.45% | -2.574% | 2.484%
| Z-Score (Absolute) |  | 0.23 | 1.81 | 0.67 | 4.5 | 0.02 | 0 | 0.55 | 1.42 | 4.09 | 1.23 | 2.89 | 0.15 | 0.4
| Analytical results for Yb in OREAS 465 (Certified Value 14.9 ppm) | [15 empty] | Certified
Yb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS |  |  | 1
2
| 1 | 1 | 14.45 | 14.45 | 14.1 | 14.65 | 14.0533 | 13.5514 | 16.6 | 15 | 10.86 | 18.8367 | 16.7 | 15 |  |  | 1
| 1 | 2 | 12.65 | 14.2 | 14 | 15 | 13.9401 | 14.3975 | 16.4 | 14.7 | 11.5 | 18.037 | 16.2 | 14.5 |  |  | 9
| 1 | 3 | 14.55 | 14.4 | 14.7 | 14.45 | 13.7687 | 14.2891 | 16.2 | 14.5 | 10.67 | 19.4474 | 16.45 | 14.8 |  |  | 16
| 1 | 4 | 14 | 16.3 | 14.65 | 14.35 | 13.6942 | 14.0654 | 16.8 | 14.2 | 10.81 | 17.911 | 16.2 | 15.2 |  |  | 14.9004
| 1 | 5 | 14.45 | 16.55 | 14.5 | 14.6 | 14.0151 | 13.9336 | 16.6 | 14.5 | 13.65 | 17.8183 | 16.25 | 15.7 |  |  | 107
| 1 | 6 | 14.75 | 14.9 | 14.55 | 14.05 | 13.7444 | 14.082 | 16.8 | 15.4 | 11.65 | 17.9023 | 16.05 | 14.7
| Mean |  | 14.1417 | 15.1333 | 14.4167 | 14.5167 | 13.8693 | 14.0532 | 16.5667 | 14.7167 | 11.5233 | 18.3254 | 16.3083 | 14.9833
| Median |  | 14.45 | 14.675 | 14.525 | 14.525 | 13.8544 | 14.0737 | 16.6 | 14.6 | 11.18 | 17.974 | 16.225 | 14.9
| Std Dev. |  | 0.77098 | 1.0294 | 0.29439 | 0.31885 | 0.15266 | 0.29705 | 0.23381 | 0.42622 | 1.115 | 0.66502 | 0.23112 | 0.42622
| Rel.Std.Dev. |  | 5.452% | 6.802% | 2.042% | 2.196% | 1.101% | 2.114% | 1.411% | 2.896% | 9.676% | 3.629% | 1.417% | 2.845%
| PDM3 |  | -5.092% | 1.563% | -3.247% | -2.575% | -6.92% | -5.686% | 11.18% | -1.233% | -22.66% | 22.99% | 9.449% | 0.5565%
| Z-Score (Absolute) |  | 0.22 | 0.65 | 0.25 | 0.12 | 0.93 | 0.7 | 2.43 | 0.12 | 4.39 | 4.35 | 2.11 | 0.46
| Analytical results for Zn in OREAS 465 (Certified Value 921 ppm) | [15 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 17 | 22 | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 934 | 1050 | 508 | 549 | 458 | 1023.5175 | 945.9699 | 877 | 926 | 1081 | 1007.6529 | 791 | 945 | 692 | 1
| 1 | 2 | 961 | 859 | 509 | 532 | 467 | 1040.4657 | 969.0924 | 890 | 941 | 1039 | 1009.8514 | 807 | 950 | 750 | 10
| 1 | 3 | 875 | 1040 | 500 | 545 | 481 | 1024.6169 | 965.3557 | 962 | 935 | 1080 | 1010.2758 | 779 | 950 | 688 | 16
| 1 | 4 | 951 | 1080 | 513 | 527 | 484 | 1038.1358 | 957.7559 | 998 | 935 | 999 | 1010.491 | 789 | 955 | 794 | 921.4187
| 1 | 5 | 898 | 1080 | 507 | 862 | 479 | 1026.3031 | 970.343 | 1032 | 936 | 1029 | 1010.724 | 775 | 950 | 796 | 108
| 1 | 6 | 881 | 1050 | 512 | 891 | 458 | 1046.9826 | 957.1705 | 901 | 932 | 1057 | 1010.4432 | 783 | 945 | 862
| Mean |  | 916.6667 | 1026.5 | 508.1667 | 651 | 471.1667 | 1033.3369 | 960.9479 | 943.3333 | 934.1667 | 1047.5 | 1009.9064 | 787.3333 | 949.1667 | 763.6667
| Median |  | 916 | 1050 | 508.5 | 547 | 473 | 1032.2194 | 961.5558 | 931.5 | 935 | 1048 | 1010.3595 | 786 | 950 | 772
| Std Dev. |  | 36.8818 | 83.7466 | 4.6224 | 175.0988 | 11.7204 | 9.8182 | 9.196 | 63.6166 | 4.9565 | 31.7348 | 1.1416 | 11.3431 | 3.7639 | 67.3548
| Rel.Std.Dev. |  | 4.023% | 8.158% | 0.9096% | 26.9% | 2.488% | 0.9501% | 0.957% | 6.744% | 0.5306% | 3.03% | 0.113% | 1.441% | 0.3965% | 8.82%
| PDM3 |  | -0.5157% | 11.4% | -44.85% | -29.35% | -48.87% | 12.15% | 4.29% | 2.378% | 1.384% | 13.68% | 9.603% | -14.55% | 3.011% | -17.12%
| Z-Score (Absolute) |  | 0.15 | 0.8 | 2.86 | 1.91 | 3.1 | 0.63 | 0.15 | 0.03 | 0.03 | 0.72 | 0.47 | 1 | 0.07 | 1.16
| Analytical results for Zr in OREAS 465 (Indicative Value 367 ppm) | [15 empty] | Indicative
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 05 | 06 | 07 | 09 | 10 | 11 | 13 | 17 | 22 |  |  | ppm
|  |  | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS |  |  | 0
0
| 1 | 1 | 486 | > 500 | > 500 | > 500 | > 500 | 545.1116 | 528 | 281 | 288.7 | 119 | 220 | FL |  |  | 1
| 1 | 2 | > 500 | > 500 | > 500 | 484 | > 500 | 525.0981 | 457 | 274 | 415.1 | 95.9 | 255 | FL |  |  | 11
| 1 | 3 | 483 | > 500 | > 500 | > 500 | > 500 | 539.7474 | 373 | 267 | 377.4 | 103 | 110 | FL |  |  | 16
| 1 | 4 | > 500 | > 500 | > 500 | > 500 | > 500 | 511.9843 | 612 | 258 | 273 | 88.9 | 155 | FL |  |  | 367.3325
| 1 | 5 | 496 | > 500 | > 500 | > 500 | > 500 | 551.6 | 630 | 277 | 499.8 | 106 | 160 | FL |  |  | 31
| 1 | 6 | 470 | > 500 | > 500 | 434 | > 500 | 514.3168 | 637 | 279 | 446.6 | 96.2 | 105 | FL
| Mean |  | 01: 483.75; 05: 459; 07: 531.3097; 08: 539.5; 09: 272.6667; 10: 383.4333; 11: 101.5; 12: 167.5
| Median |  | 01: 484.5; 05: 459; 07: 532.4228; 08: 570; 09: 275.5; 10: 396.25; 11: 99.6; 12: 157.5
| Std Dev. |  | 01: 10.7199; 05: 35.3553; 07: 16.5783; 08: 107.3438; 09: 8.6872; 10: 89.1519; 11: 10.4591; 12: 59.7286
| Rel.Std.Dev. |  | 01: 2.216%; 05: 7.703%; 07: 3.12%; 08: 19.9%; 09: 3.186%; 10: 23.25%; 11: 10.3%; 12: 35.66%
| PDM3 |  | 01: 31.69%; 05: 24.95%; 07: 44.64%; 08: 46.87%; 09: -25.77%; 10: 4.383%; 11: -72.37%; 12: -54.4%
| Z-Score (Absolute) |  | 0.37 | NA | NA | 0.22 | NA | 0.65 | 0.7 | 0.88 | 0.22 | 1.89 | 1.5 | NA
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.