## Abbreviations
Table 1. Abbreviations used for OREAS 905
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one sigma relative standard deviation
Std.Dev. | one sigma standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
ABL*MS | Laser Ablation with Mass Spectrometry: ICP-MS finish
AL-5S*AAS | acid leach of 5% sulphuric acid with AAS finish
AL-5S*OES | acid leach of 5% sulphuric acid with OES finish
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
at 1000°C | loss on ignition at 1000°C
BF*MS | lithium borate fusion with ICP-MS finish
BF*OES | lithium borate fusion with ICP-OES finish
BF*XRF | lithium borate fusion with XRF finish
FA*AAS | fire assay with AAS finish
FA*MS | fire assay with ICP-MS finish
FA*NAA | fire assay with neutron activation analysis finish
FA*OES | fire assay with ICP-OES finish
GASPYC | gas expansion pycnometer
INAA | INAA using a charge weight as deemed appropriate
IRC | infrared combustion furnace
LIQPYC | Liquid pycnometry
PF*AAS | sodium peroxide fusion with AAS finish
PF*MS | sodium peroxide fusion with ICP-MS finish
PF*OES | sodium peroxide fusion with ICP-OES finish
PF*OES/MS | sodium peroxide fusion with OES or MS finish as appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 905
Lab Slot | Laboratory Name and Location
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Johannesburg, South Africa
ALS, Lima, Peru
ALS, Loughrea, Galway, Ireland
ALS, Perth, WA, Australia
ALS, Reno, Nevada, USA
ALS, Vancouver, BC, Canada
American Assay Laboratories, Sparks, Nevada, USA
ANSTO, Lucas Heights, NSW, Australia
Bureau Veritas Commodities Canada Ltd, Vancouver, BC, Canada
Bureau Veritas Geoanalytical, Adelaide, SA, Australia
Bureau Veritas Geoanalytical, Perth, WA, Australia
Bureau Veritas Minerals, Santiago, Chile
CIMM TyS S.A., Antofagasta, Chile
Inspectorate (BV), Lima, Peru
Inspectorate America Corporation (BV), Sparks, Nevada, USA
Intertek Genalysis, Adelaide, SA, Australia
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Shunyi, Beijing, China
McClelland Laboratories Inc., Sparks, Nevada, USA
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
PT Intertek Utama Services, Jakarta Timur, DKI Jakarta, Indonesia
SGS, Sudbury, Ontario, Canada
SGS Australia Mineral Services, Perth (Newburn), WA, Australia
SGS Canada Inc., Vancouver, BC, Canada
SGS del Peru, Lima, Peru
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Lakefield Research Ltd, Lakefield, Ontario, Canada
SGS Mineral Services, Townsville, QLD, Australia
SGS Minerals, Santiago, Chile
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
Skyline, Sparks, Nevada, USA
TSL Laboratories Inc., Saskatoon, Saskatchewan, Canada

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 905
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.39081 | 0.0093584 | 0.38762 | 0.39399 | 0.3878 | 0.39381
4-Acid Digestion
Ag, Silver (ppm) | 0.51764 | 0.095101 | 0.47761 | 0.55766 | 0.48917 | 0.54611
Al, Aluminium (wt.%) | 7.4245 | 0.26551 | 7.3285 | 7.5204 | 7.2182 | 7.6308
As, Arsenic (ppm) | 34.6906 | 1.7398 | 34.0606 | 35.3206 | 33.3942 | 35.987
Ba, Barium (ppm) | 2699.3903 | 125.9364 | 2652.0806 | 2746.7 | 2651.4242 | 2747.3564
Be, Beryllium (ppm) | 3.0438 | 0.3167 | 2.9325 | 3.1552 | 2.9056 | 3.1821
Bi, Bismuth (ppm) | 5.7152 | 0.33688 | 5.5574 | 5.873 | 5.5677 | 5.8627
Ca, Calcium (wt.%) | 0.59011 | 0.027853 | 0.5803 | 0.59991 | 0.5753 | 0.60492
Cd, Cadmium (ppm) | 0.35727 | 0.044843 | 0.33711 | 0.37743 | 0.32801 | 0.38654
Ce, Cerium (ppm) | 92.1763 | 4.3131 | 90.2716 | 94.0811 | 89.151 | 95.2016
Co, Cobalt (ppm) | 14.8446 | 0.80434 | 14.5472 | 15.142 | 14.4519 | 15.2373
Cr, Chromium (ppm) | 19.1784 | 3.1401 | 17.9264 | 20.4305 | 17.449 | 20.9079
Cs, Cesium (ppm) | 6.7784 | 0.40716 | 6.5986 | 6.9583 | 6.5971 | 6.9598
Cu, Copper (ppm) | 1533.4576 | 60.6345 | 1513.0404 | 1553.8748 | 1497.5475 | 1569.3677
Dy, Dysprosium (ppm) | 3.7245 | 0.21645 | 3.5816 | 3.8675 | 3.5296 | 3.9194
Er, Erbium (ppm) | 1.1164 | 0.045683 | 1.0905 | 1.1424 | 1.0467 | 1.1861
Eu, Europium (ppm) | 1.4225 | 0.11473 | 1.3372 | 1.5078 | 1.3263 | 1.5188
Fe, Iron (wt.%) | 4.0764 | 0.15618 | 4.0201 | 4.1327 | 3.9749 | 4.1779
Ga, Gallium (ppm) | 25.118 | 1.0176 | 24.6917 | 25.5443 | 24.2617 | 25.9743
Gd, Gadolinium (ppm) | 5.8982 | 0.36393 | 5.6023 | 6.1942 | 5.5808 | 6.2157
Hf, Hafnium (ppm) | 6.8444 | 0.37748 | 6.6827 | 7.0061 | 6.6142 | 7.0746
Ho, Holmium (ppm) | 0.4987 | 0.046066 | 0.47219 | 0.52521 | 0.46682 | 0.53059
In, Indium (ppm) | 0.63586 | 0.047292 | 0.61512 | 0.6566 | 0.61279 | 0.65893
K, Potassium (wt.%) | 2.8752 | 0.11589 | 2.837 | 2.9134 | 2.7969 | 2.9535
La, Lanthanum (ppm) | 45.9619 | 2.9687 | 44.8195 | 47.1044 | 44.4334 | 47.4904
Li, Lithium (ppm) | 20.0438 | 1.5998 | 19.4241 | 20.6635 | 19.3905 | 20.6971
Lu, Lutetium (ppm) | 0.10111 | 0.014631 | 0.090566 | 0.11166 | IND | IND
Mg, Magnesium (wt.%) | 0.27622 | 0.015854 | 0.27052 | 0.28192 | 0.26742 | 0.28502
Mn, Manganese (wt.%) | 0.037747 | 0.0016315 | 0.037163 | 0.03833 | 0.036836 | 0.038658
Mo, Molybdenum (ppm) | 3.2703 | 0.26172 | 3.1863 | 3.3543 | 3.1225 | 3.4181
Na, Sodium (wt.%) | 2.3997 | 0.1124 | 2.3577 | 2.4417 | 2.3486 | 2.4508
Nb, Niobium (ppm) | 18.1385 | 1.0973 | 17.7114 | 18.5657 | 17.4695 | 18.8076
Nd, Neodymium (ppm) | 39.0542 | 1.5464 | 38.3014 | 39.807 | 37.5867 | 40.5217
Ni, Nickel (ppm) | 9.5446 | 0.84432 | 9.2744 | 9.8148 | 8.9692 | 10.12
P, Phosphorus (wt.%) | 0.027505 | 0.0022517 | 0.026664 | 0.028346 | 0.026443 | 0.028566
Pb, Lead (ppm) | 30.4419 | 2.0979 | 29.6503 | 31.2334 | 29.339 | 31.5447
Pr, Praseodymium (ppm) | 10.5396 | 0.63626 | 10.0982 | 10.9809 | 10.2023 | 10.8768
Rb, Rubidium (ppm) | 138.2814 | 6.3663 | 135.555 | 141.0078 | 134.3348 | 142.2279
S, Sulphur (wt.%) | 0.066039 | 0.0064718 | 0.063624 | 0.068455 | 0.063201 | 0.068878
Sb, Antimony (ppm) | 1.9516 | 0.12815 | 1.898 | 2.0051 | 1.8485 | 2.0547
Sc, Scandium (ppm) | 4.8952 | 0.45935 | 4.7281 | 5.0624 | 4.7052 | 5.0852
Se, Selenium (ppm) | 2.8408 | 0.32219 | 2.7192 | 2.9624 | 2.1337 | 3.5479
Sm, Samarium (ppm) | 7.3993 | 0.50639 | 7.0428 | 7.7558 | 7.0759 | 7.7228
Sn, Tin (ppm) | 3.9563 | 0.27071 | 3.8546 | 4.0579 | 3.7797 | 4.1329
Sr, Strontium (ppm) | 156.9475 | 6.5151 | 154.6233 | 159.2717 | 152.3392 | 161.5558
Ta, Tantalum (ppm) | 1.345 | 0.082862 | 1.3075 | 1.3824 | 1.2536 | 1.4364
Tb, Terbium (ppm) | 0.7703 | 0.069586 | 0.72778 | 0.81282 | 0.73826 | 0.80234
Te, Tellurium (ppm) | 0.0878 | 0.017063 | 0.078881 | 0.096719 | IND | IND
Th, Thorium (ppm) | 14.6454 | 0.75992 | 14.3494 | 14.9413 | 14.1881 | 15.1026
Ti, Titanium (wt.%) | 0.12192 | 0.006241 | 0.11973 | 0.1241 | 0.11903 | 0.12481
Tl, Thallium (ppm) | 0.72195 | 0.055024 | 0.69819 | 0.74571 | 0.69461 | 0.74929
Tm, Thulium (ppm) | 0.10889 | 0.013991 | 0.098325 | 0.11945 | IND | IND
U, Uranium (ppm) | 4.9743 | 0.30587 | 4.8433 | 5.1052 | 4.828 | 5.1205
V, Vanadium (ppm) | 10.1466 | 0.96736 | 9.7794 | 10.5138 | IND | IND
W, Tungsten (ppm) | 2.7768 | 0.17381 | 2.712 | 2.8416 | 2.6033 | 2.9504
Y, Yttrium (ppm) | 15.7134 | 0.80646 | 15.4377 | 15.989 | 15.2676 | 16.1591
Yb, Ytterbium (ppm) | 0.675 | 0.047602 | 0.64319 | 0.70681 | IND | IND
Zn, Zinc (ppm) | 137.9507 | 6.6595 | 135.6079 | 140.2934 | 133.6861 | 142.2152
Zr, Zirconium (ppm) | 252.1642 | 11.5551 | 247.8627 | 256.4658 | 246.3253 | 258.0032
Borate / Peroxide Fusion ICP
Al, Aluminium (wt.%) | 7.5948 | 0.17451 | 7.526 | 7.6636 | 7.4272 | 7.7624
As, Arsenic (ppm) | 36.1983 | 4.6495 | 34.0327 | 38.3639 | 32.6572 | 39.7395
Ba, Barium (ppm) | 2795.3371 | 105.4601 | 2740.4426 | 2850.2317 | 2731.5774 | 2859.0969
Be, Beryllium (ppm) | 3.0828 | 0.52274 | 2.7512 | 3.4143 | 2.9111 | 3.2544
Bi, Bismuth (ppm) | 6.1971 | 0.35724 | 5.9522 | 6.4421 | 5.7214 | 6.6729
Ca, Calcium (wt.%) | 0.60822 | 0.028674 | 0.59642 | 0.62003 | 0.57622 | 0.64022
Ce, Cerium (ppm) | 95.8806 | 5.9077 | 92.963 | 98.7982 | 91.8133 | 99.948
Co, Cobalt (ppm) | 15.2807 | 2.1279 | 14.6493 | 15.912 | 14.7574 | 15.804
Cs, Cesium (ppm) | 7.0989 | 0.43916 | 6.8467 | 7.3511 | 6.8435 | 7.3542
Cu, Copper (ppm) | 1571.4521 | 52.711 | 1554.2765 | 1588.6276 | 1526.9216 | 1615.9826
Dy, Dysprosium (ppm) | 3.7842 | 0.34019 | 3.5504 | 4.018 | 3.5698 | 3.9985
Er, Erbium (ppm) | 1.1377 | 0.11657 | 1.0505 | 1.2249 | 1.0225 | 1.2528
Eu, Europium (ppm) | 1.4579 | 0.088585 | 1.4039 | 1.5118 | 1.3541 | 1.5616
Fe, Iron (wt.%) | 4.2268 | 0.094096 | 4.1904 | 4.2632 | 4.1331 | 4.3205
Ga, Gallium (ppm) | 25.8159 | 1.4196 | 25.0834 | 26.5485 | 24.6736 | 26.9583
Gd, Gadolinium (ppm) | 6.4258 | 0.74536 | 5.8551 | 6.9965 | 6.1349 | 6.7167
Hf, Hafnium (ppm) | 7.2583 | 0.7373 | 6.6658 | 7.8508 | 6.9045 | 7.6121
Ho, Holmium (ppm) | 0.53431 | 0.070188 | 0.47942 | 0.5892 | 0.45043 | 0.61819
In, Indium (ppm) | 0.62198 | 0.076009 | 0.57238 | 0.67159 | IND | IND
K, Potassium (wt.%) | 2.9394 | 0.12246 | 2.8876 | 2.9913 | 2.8595 | 3.0194
La, Lanthanum (ppm) | 47.9702 | 2.7169 | 46.6673 | 49.2731 | 45.8186 | 50.1218
Li, Lithium (ppm) | 20.6086 | 3.1081 | 18.6568 | 22.5603 | 19.0609 | 22.1562
Mg, Magnesium (wt.%) | 0.28787 | 0.015173 | 0.28143 | 0.2943 | 0.27751 | 0.29823
Mn, Manganese (wt.%) | 0.039483 | 0.001253 | 0.039011 | 0.039955 | 0.038105 | 0.040862
Nb, Niobium (ppm) | 18.5042 | 2.7194 | 16.8934 | 20.1151 | 16.0666 | 20.9419
Nd, Neodymium (ppm) | 40.4846 | 1.8185 | 39.4205 | 41.5488 | 37.8223 | 43.1469
P, Phosphorus (wt.%) | 0.02829 | 0.002871 | 0.026965 | 0.029616 | 0.025668 | 0.030912
Pb, Lead (ppm) | 30.8094 | 3.3005 | 28.5229 | 33.0959 | 28.2522 | 33.3666
Pr, Praseodymium (ppm) | 10.8876 | 0.73128 | 10.3848 | 11.3905 | 10.4157 | 11.3596
Rb, Rubidium (ppm) | 137.4463 | 4.5166 | 134.6194 | 140.2733 | 133.7264 | 141.1662
S, Sulphur (wt.%) | 0.071317 | 0.013132 | 0.066004 | 0.076629 | IND | IND
Sb, Antimony (ppm) | 1.9566 | 0.32233 | 1.7692 | 2.1441 | 1.6086 | 2.3046
Si, Silicon (wt.%) | 32.3295 | 1.1333 | 31.6807 | 32.9782 | 31.7601 | 32.8988
Sm, Samarium (ppm) | 7.6355 | 0.28611 | 7.5059 | 7.7652 | 7.2522 | 8.0189
Sn, Tin (ppm) | 7.4879 | 1.1858 | 6.1154 | 8.8604 | 6.2483 | 8.7275
Sr, Strontium (ppm) | 158.6086 | 6.4929 | 155.1559 | 162.0612 | 148.8796 | 168.3376
Ta, Tantalum (ppm) | 1.376 | 0.14387 | 1.2851 | 1.4668 | 1.1717 | 1.5802
Tb, Terbium (ppm) | 0.81381 | 0.10046 | 0.73797 | 0.88966 | 0.70937 | 0.91826
Th, Thorium (ppm) | 14.6884 | 0.64055 | 14.3103 | 15.0665 | 14.0949 | 15.2819
Ti, Titanium (wt.%) | 0.12759 | 0.0050742 | 0.12548 | 0.12969 | 0.12346 | 0.13171
Tl, Thallium (ppm) | 0.7437 | 0.04362 | 0.71919 | 0.76822 | IND | IND
Tm, Thulium (ppm) | 0.14077 | 0.025741 | 0.13123 | 0.15031 | IND | IND
U, Uranium (ppm) | 5.0041 | 0.4113 | 4.7195 | 5.2887 | 4.7822 | 5.226
V, Vanadium (ppm) | 11.2958 | 1.5717 | 10.2021 | 12.3896 | IND | IND
W, Tungsten (ppm) | 3.0182 | 0.45623 | 2.8439 | 3.1925 | 2.5694 | 3.467
Y, Yttrium (ppm) | 16.5676 | 1.4194 | 15.8732 | 17.262 | 16.0595 | 17.0757
Yb, Ytterbium (ppm) | 0.75552 | 0.072759 | 0.69961 | 0.81143 | IND | IND
Zn, Zinc (ppm) | 138.9064 | 9.9651 | 134.1509 | 143.662 | 131.7243 | 146.0886
Zr, Zirconium (ppm) | 270.0568 | 18.6084 | 251.2254 | 288.8882 | 259.7811 | 280.3325
Aqua Regia Digestion
Ag, Silver (ppm) | 0.51648 | 0.048596 | 0.49935 | 0.53361 | 0.48874 | 0.54422
Al, Aluminium (wt.%) | 0.81726 | 0.049255 | 0.79681 | 0.8377 | 0.79594 | 0.83858
As, Arsenic (ppm) | 31.6861 | 1.8173 | 31.0069 | 32.3652 | 30.487 | 32.8852
Au, Gold (ppm) | 0.39498 | 0.018885 | 0.38672 | 0.40324 | 0.39166 | 0.39831
Ba, Barium (ppm) | 248.984 | 15.4793 | 242.9079 | 255.0601 | 242.1327 | 255.8353
Be, Beryllium (ppm) | 0.93372 | 0.079028 | 0.90398 | 0.96347 | 0.8673 | 1.0001
Bi, Bismuth (ppm) | 5.7447 | 0.35304 | 5.5905 | 5.899 | 5.5495 | 5.9399
Ca, Calcium (wt.%) | 0.33678 | 0.017764 | 0.32995 | 0.34361 | 0.32718 | 0.34637
Cd, Cadmium (ppm) | 0.34212 | 0.02872 | 0.33241 | 0.35182 | 0.32522 | 0.35901
Ce, Cerium (ppm) | 79.5606 | 5.2544 | 77.2775 | 81.8437 | 77.1806 | 81.9405
Co, Cobalt (ppm) | 13.8923 | 0.90601 | 13.5728 | 14.2119 | 13.4477 | 14.337
Cr, Chromium (ppm) | 17.6185 | 1.813 | 16.9125 | 18.3246 | 16.2278 | 19.0093
Cs, Cesium (ppm) | 1.3175 | 0.16352 | 1.2387 | 1.3963 | 1.2522 | 1.3829
Cu, Copper (ppm) | 1561.789 | 47.1271 | 1545.267 | 1578.3111 | 1533.7289 | 1589.8492
Dy, Dysprosium (ppm) | 1.7302 | 0.32432 | 1.4737 | 1.9866 | 1.6313 | 1.8291
Eu, Europium (ppm) | 0.91738 | 0.091507 | 0.84031 | 0.99444 | 0.88744 | 0.94731
Fe, Iron (wt.%) | 3.4952 | 0.14925 | 3.4357 | 3.5546 | 3.4297 | 3.5606
Ga, Gallium (ppm) | 6.4321 | 0.66058 | 6.1506 | 6.7136 | 6.2221 | 6.6421
Gd, Gadolinium (ppm) | 3.5854 | 0.34381 | 3.3032 | 3.8676 | 3.4322 | 3.7387
Ge, Germanium (ppm) | < 1 | IND | IND | IND | IND | IND
Hf, Hafnium (ppm) | 1.2219 | 0.22161 | 1.1019 | 1.342 | 1.1663 | 1.2776
Hg, Mercury (ppm) | < 0.05 | IND | IND | IND | IND | IND
Ho, Holmium (ppm) | 0.22258 | 0.036577 | 0.19027 | 0.2549 | 0.21262 | 0.23255
In, Indium (ppm) | 0.5788 | 0.042398 | 0.56011 | 0.5975 | 0.55002 | 0.60759
K, Potassium (wt.%) | 0.32201 | 0.023125 | 0.31271 | 0.33132 | 0.31231 | 0.33171
La, Lanthanum (ppm) | 39.8025 | 2.4565 | 38.8883 | 40.7167 | 38.7046 | 40.9004
Li, Lithium (ppm) | 4.8746 | 0.66847 | 4.5948 | 5.1543 | 4.6395 | 5.1096
Lu, Lutetium (ppm) | 0.033944 | 0.0053262 | 0.030645 | 0.037243 | IND | IND
Mg, Magnesium (wt.%) | 0.15831 | 0.011861 | 0.15364 | 0.16298 | 0.15352 | 0.1631
Mn, Manganese (wt.%) | 0.03528 | 0.0011627 | 0.034857 | 0.035703 | 0.034436 | 0.036124
Mo, Molybdenum (ppm) | 3.0239 | 0.17285 | 2.9723 | 3.0755 | 2.9026 | 3.1452
Na, Sodium (wt.%) | 0.092062 | 0.0093871 | 0.088641 | 0.095484 | 0.088616 | 0.095508
Nd, Neodymium (ppm) | 30.4811 | 3.5143 | 27.7169 | 33.2453 | 29.4555 | 31.5067
Ni, Nickel (ppm) | 8.903 | 0.6374 | 8.6932 | 9.1127 | 8.2925 | 9.5134
P, Phosphorus (wt.%) | 0.023418 | 0.0016927 | 0.02272 | 0.024116 | 0.022522 | 0.024314
Pb, Lead (ppm) | 17.0849 | 1.6681 | 16.5207 | 17.649 | 16.4302 | 17.7396
Pr, Praseodymium (ppm) | 8.4199 | 0.69814 | 7.9001 | 8.9396 | 8.1039 | 8.7358
Rb, Rubidium (ppm) | 19.2514 | 1.4096 | 18.6001 | 19.9026 | 18.4139 | 20.0888
S, Sulphur (wt.%) | 0.065101 | 0.0065831 | 0.062687 | 0.067516 | 0.062439 | 0.067764
Sb, Antimony (ppm) | 1.1247 | 0.16323 | 1.0497 | 1.1996 | 1.0668 | 1.1825
Sc, Scandium (ppm) | 1.8901 | 0.19437 | 1.8127 | 1.9674 | 1.6926 | 2.0875
Se, Selenium (ppm) | 2.3169 | 0.43504 | 2.123 | 2.5108 | 2.0811 | 2.5527
Sm, Samarium (ppm) | 5.0524 | 0.45465 | 4.633 | 5.4718 | 4.88 | 5.2247
Sn, Tin (ppm) | 1.2715 | 0.1375 | 1.2092 | 1.3337 | 1.1282 | 1.4148
Sr, Strontium (ppm) | 12.3315 | 1.1698 | 11.9073 | 12.7556 | 11.7504 | 12.9125
Tb, Terbium (ppm) | 0.43862 | 0.052166 | 0.40337 | 0.47387 | 0.42214 | 0.4551
Te, Tellurium (ppm) | 0.065111 | 0.011375 | 0.061435 | 0.068788 | IND | IND
Th, Thorium (ppm) | 8.8816 | 0.70476 | 8.5651 | 9.1982 | 8.5953 | 9.168
Ti, Titanium (wt.%) | 0.019173 | 0.0027197 | 0.018067 | 0.02028 | 0.018109 | 0.020238
Tl, Thallium (ppm) | 0.10972 | 0.011326 | 0.10559 | 0.11386 | IND | IND
Tm, Thulium (ppm) | 0.0473 | 0.0078017 | 0.0378 | 0.056801 | IND | IND
U, Uranium (ppm) | 2.3721 | 0.19441 | 2.2851 | 2.4591 | 2.2737 | 2.4705
V, Vanadium (ppm) | 6.0549 | 0.84668 | 5.7374 | 6.3724 | 5.7827 | 6.3271
W, Tungsten (ppm) | < 1 | IND | IND | IND | IND | IND
Y, Yttrium (ppm) | 7.0814 | 0.49223 | 6.8912 | 7.2715 | 6.7769 | 7.3858
Yb, Ytterbium (ppm) | 0.26733 | 0.040945 | 0.2467 | 0.28797 | 0.23717 | 0.2975
Zn, Zinc (ppm) | 67.1665 | 5.0341 | 65.1661 | 69.1668 | 64.4116 | 69.9214
Zr, Zirconium (ppm) | 47.4967 | 4.9543 | 45.1817 | 49.8117 | 45.6362 | 49.3572
Sulphuric Acid Leach
Cu-Sol, Copper Soluble (ppm) | 1271.951 | 65.4312 | 1245.3967 | 1298.5053 | 1249.4043 | 1294.4978
Gas / Liquid Pycnometry
SG, Specific Gravity (Unity) | 2.7473 | 0.04168 | 2.7273 | 2.7673 | 2.727 | 2.7676

## Indicative Values
Table 3. Indicative Values for OREAS 905
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | < 1 | Pt | ppb | < 1
4-Acid Digestion
B | ppm | 4 | Hg | ppm | 0.026667
Ge | ppm | 0.53686 | Re | ppm | < 0.002
Borate / Peroxide Fusion ICP
Ag | ppm | 0.685 | Lu | ppm | 0.083733 | Sc | ppm | 4.4997
B | ppm | 19.1667 | Mo | ppm | 2.9487 | Se | ppm | < 20
Cd | ppm | 0.4025 | Na | wt.% | 2.4075 | Te | ppm | < 1
Cr | ppm | 28.2183 | Ni | ppm | 12.5381
Ge | ppm | 1.5303 | Re | ppm | < 0.1
Aqua Regia Digestion
B | ppm | < 10 | Os | ppm | < 0.01 | Rh | ppm | 0.065
Er | ppm | 0.44082 | Pd | ppb | < 10 | Ru | ppm | 2.6783
Ir | ppm | < 0.01 | Pt | ppb | < 5 | Ta | ppm | < 0.01
Nb | ppm | 0.43926 | Re | ppm | < 0.001
Borate Fusion XRF
Al2O3 | wt.% | 14.5667 | MgO | wt.% | 0.49667 | SiO2 | wt.% | 68.25
CaO | wt.% | 0.85833 | MnO | wt.% | 0.05 | SO3 | wt.% | 0.17
Fe2O3 | wt.% | 5.8833 | Na2O | wt.% | 3.395 | TiO2 | wt.% | 0.21
K2O | wt.% | 3.5717 | P2O5 | wt.% | 0.0685
Thermogravimetry
LOI1000 | wt.% | 1.8558
Infrared Combustion
S | wt.% | 0.053333

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 905
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.39081 | 0.0093584 | 0.37209 | 0.40952 | 0.36273 | 0.41888 | 2.395% | 4.789% | 7.184% | 0.37127 | 0.41035
4-Acid Digestion
Ag, ppm | 0.51764 | 0.095101 | 0.32744 | 0.70784 | 0.23234 | 0.80294 | 18.37% | 36.74% | 55.12% | 0.49176 | 0.54352
Al, wt.% | 7.4245 | 0.26551 | 6.8935 | 7.9555 | 6.6279 | 8.221 | 3.576% | 7.152% | 10.73% | 7.0533 | 7.7957
As, ppm | 34.6906 | 1.7398 | 31.211 | 38.1702 | 29.4712 | 39.91 | 5.015% | 10.03% | 15.05% | 32.9561 | 36.4251
Ba, ppm | 2699.3903 | 125.9364 | 2447.5174 | 2951.2632 | 2321.581 | 3077.1996 | 4.665% | 9.331% | 14% | 2564.4208 | 2834.3598
Be, ppm | 3.0438 | 0.3167 | 2.4104 | 3.6772 | 2.0937 | 3.9939 | 10.4% | 20.81% | 31.21% | 2.8917 | 3.196
Bi, ppm | 5.7152 | 0.33688 | 5.0414 | 6.3889 | 4.7046 | 6.7258 | 5.894% | 11.79% | 17.68% | 5.4294 | 6.0009
Ca, wt.% | 0.59011 | 0.027853 | 0.5344 | 0.64581 | 0.50655 | 0.67367 | 4.72% | 9.44% | 14.16% | 0.5606 | 0.61961
Cd, ppm | 0.35727 | 0.044843 | 0.26759 | 0.44696 | 0.22274 | 0.4918 | 12.55% | 25.1% | 37.65% | 0.33941 | 0.37514
Ce, ppm | 92.1763 | 4.3131 | 83.5501 | 100.8026 | 79.237 | 105.1157 | 4.679% | 9.358% | 14.04% | 87.5675 | 96.7852
Co, ppm | 14.8446 | 0.80434 | 13.2359 | 16.4533 | 12.4316 | 17.2576 | 5.418% | 10.84% | 16.26% | 14.1024 | 15.5868
Cr, ppm | 19.1784 | 3.1401 | 12.8983 | 25.4586 | 9.7582 | 28.5987 | 16.37% | 32.75% | 49.12% | 18.2195 | 20.1374
Cs, ppm | 6.7784 | 0.40716 | 5.9641 | 7.5927 | 5.557 | 7.9999 | 6.007% | 12.01% | 18.02% | 6.4395 | 7.1174
Cu, ppm | 1533.4576 | 60.6345 | 1412.1886 | 1654.7266 | 1351.5541 | 1715.361 | 3.954% | 7.908% | 11.86% | 1456.7847 | 1610.1305
Dy, ppm | 3.7245 | 0.21645 | 3.2916 | 4.1574 | 3.0752 | 4.3739 | 5.812% | 11.62% | 17.43% | 3.5383 | 3.9108
Er, ppm | 1.1164 | 0.045683 | 1.0251 | 1.2078 | 0.97938 | 1.2535 | 4.092% | 8.184% | 12.28% | 1.0606 | 1.1723
Eu, ppm | 1.4225 | 0.11473 | 1.1931 | 1.652 | 1.0783 | 1.7667 | 8.065% | 16.13% | 24.2% | 1.3514 | 1.4937
Fe, wt.% | 4.0764 | 0.15618 | 3.7641 | 4.3888 | 3.6079 | 4.5449 | 3.831% | 7.663% | 11.49% | 3.8726 | 4.2802
Ga, ppm | 25.118 | 1.0176 | 23.0828 | 27.1532 | 22.0652 | 28.1708 | 4.051% | 8.103% | 12.15% | 23.8621 | 26.3739
Gd, ppm | 5.8982 | 0.36393 | 5.1704 | 6.6261 | 4.8065 | 6.99 | 6.17% | 12.34% | 18.51% | 5.6033 | 6.1932
Hf, ppm | 6.8444 | 0.37748 | 6.0894 | 7.5994 | 5.712 | 7.9769 | 5.515% | 11.03% | 16.55% | 6.5022 | 7.1866
Ho, ppm | 0.4987 | 0.046066 | 0.40657 | 0.59084 | 0.36051 | 0.6369 | 9.237% | 18.47% | 27.71% | 0.47377 | 0.52364
In, ppm | 0.63586 | 0.047292 | 0.54127 | 0.73044 | 0.49398 | 0.77773 | 7.437% | 14.87% | 22.31% | 0.60407 | 0.66765
K, wt.% | 2.8752 | 0.11589 | 2.6434 | 3.107 | 2.5275 | 3.2229 | 4.031% | 8.062% | 12.09% | 2.7315 | 3.019
La, ppm | 45.9619 | 2.9687 | 40.0245 | 51.8994 | 37.0557 | 54.8681 | 6.459% | 12.92% | 19.38% | 43.6638 | 48.26
Li, ppm | 20.0438 | 1.5998 | 16.8443 | 23.2433 | 15.2445 | 24.8431 | 7.981% | 15.96% | 23.94% | 19.0416 | 21.046
Lu, ppm | 0.10111 | 0.014631 | 0.071848 | 0.13037 | 0.057217 | 0.14501 | 14.47% | 28.94% | 43.41% | 0.096056 | 0.10617
Mg, wt.% | 0.27622 | 0.015854 | 0.24451 | 0.30793 | 0.22866 | 0.32378 | 5.739% | 11.48% | 17.22% | 0.26241 | 0.29003
Mn, wt.% | 0.037747 | 0.0016315 | 0.034484 | 0.04101 | 0.032852 | 0.042641 | 4.322% | 8.644% | 12.97% | 0.035859 | 0.039634
Mo, ppm | 3.2703 | 0.26172 | 2.7469 | 3.7938 | 2.4852 | 4.0555 | 8.003% | 16.01% | 24.01% | 3.1068 | 3.4338
Na, wt.% | 2.3997 | 0.1124 | 2.1749 | 2.6245 | 2.0625 | 2.7369 | 4.684% | 9.368% | 14.05% | 2.2797 | 2.5197
Nb, ppm | 18.1385 | 1.0973 | 15.944 | 20.3331 | 14.8467 | 21.4304 | 6.049% | 12.1% | 18.15% | 17.2316 | 19.0455
Nd, ppm | 39.0542 | 1.5464 | 35.9615 | 42.1469 | 34.4151 | 43.6932 | 3.96% | 7.919% | 11.88% | 37.1015 | 41.0069
Ni, ppm | 9.5446 | 0.84432 | 7.856 | 11.2333 | 7.0117 | 12.0776 | 8.846% | 17.69% | 26.54% | 9.0674 | 10.0218
P, wt.% | 0.027505 | 0.0022517 | 0.023001 | 0.032008 | 0.020749 | 0.03426 | 8.187% | 16.37% | 24.56% | 0.026129 | 0.02888
Pb, ppm | 30.4419 | 2.0979 | 26.246 | 34.6377 | 24.148 | 36.7357 | 6.892% | 13.78% | 20.67% | 28.9198 | 31.9639
Pr, ppm | 10.5396 | 0.63626 | 9.267 | 11.8121 | 8.6308 | 12.4483 | 6.037% | 12.07% | 18.11% | 10.0126 | 11.0665
Rb, ppm | 138.2814 | 6.3663 | 125.5488 | 151.014 | 119.1825 | 157.3803 | 4.604% | 9.208% | 13.81% | 131.3673 | 145.1954
S, wt.% | 0.066039 | 0.0064718 | 0.053096 | 0.078983 | 0.046624 | 0.085455 | 9.8% | 19.6% | 29.4% | 0.062737 | 0.069341
Sb, ppm | 1.9516 | 0.12815 | 1.6953 | 2.2079 | 1.5671 | 2.336 | 6.567% | 13.13% | 19.7% | 1.854 | 2.0491
Sc, ppm | 4.8952 | 0.45935 | 3.9765 | 5.8139 | 3.5172 | 6.2733 | 9.384% | 18.77% | 28.15% | 4.6505 | 5.14
Se, ppm | 2.8408 | 0.32219 | 2.1964 | 3.4852 | 1.8742 | 3.8074 | 11.34% | 22.68% | 34.02% | 2.6988 | 2.9828
Sm, ppm | 7.3993 | 0.50639 | 6.3865 | 8.4121 | 5.8802 | 8.9185 | 6.844% | 13.69% | 20.53% | 7.0294 | 7.7693
Sn, ppm | 3.9563 | 0.27071 | 3.4148 | 4.4977 | 3.1441 | 4.7684 | 6.843% | 13.69% | 20.53% | 3.7584 | 4.1541
Sr, ppm | 156.9475 | 6.5151 | 143.9174 | 169.9776 | 137.4023 | 176.4926 | 4.151% | 8.302% | 12.45% | 149.1001 | 164.7949
Ta, ppm | 1.345 | 0.082862 | 1.1793 | 1.5107 | 1.0964 | 1.5936 | 6.161% | 12.32% | 18.48% | 1.2777 | 1.4122
Tb, ppm | 0.7703 | 0.069586 | 0.63113 | 0.90947 | 0.56155 | 0.97906 | 9.034% | 18.07% | 27.1% | 0.73179 | 0.80882
Te, ppm | 0.0878 | 0.017063 | 0.053673 | 0.12193 | 0.03661 | 0.13899 | 19.43% | 38.87% | 58.3% | 0.08341 | 0.09219
Th, ppm | 14.6454 | 0.75992 | 13.1255 | 16.1652 | 12.3656 | 16.9251 | 5.189% | 10.38% | 15.57% | 13.9131 | 15.3776
Ti, wt.% | 0.12192 | 0.006241 | 0.10944 | 0.1344 | 0.10319 | 0.14064 | 5.119% | 10.24% | 15.36% | 0.11582 | 0.12801
Tl, ppm | 0.72195 | 0.055024 | 0.6119 | 0.832 | 0.55688 | 0.88702 | 7.622% | 15.24% | 22.86% | 0.68585 | 0.75805
Tm, ppm | 0.10889 | 0.013991 | 0.080907 | 0.13687 | 0.066916 | 0.15086 | 12.85% | 25.7% | 38.55% | 0.10344 | 0.11433
U, ppm | 4.9743 | 0.30587 | 4.3625 | 5.586 | 4.0567 | 5.8919 | 6.149% | 12.3% | 18.45% | 4.7256 | 5.223
V, ppm | 10.1466 | 0.96736 | 8.2119 | 12.0813 | 7.2445 | 13.0487 | 9.534% | 19.07% | 28.6% | 9.6393 | 10.654
W, ppm | 2.7768 | 0.17381 | 2.4292 | 3.1244 | 2.2554 | 3.2982 | 6.259% | 12.52% | 18.78% | 2.638 | 2.9157
Y, ppm | 15.7134 | 0.80646 | 14.1004 | 17.3263 | 13.294 | 18.1328 | 5.132% | 10.26% | 15.4% | 14.9277 | 16.499
Yb, ppm | 0.675 | 0.047602 | 0.5798 | 0.7702 | 0.53219 | 0.81781 | 7.052% | 14.1% | 21.16% | 0.64125 | 0.70875
Zn, ppm | 137.9507 | 6.6595 | 124.6317 | 151.2696 | 117.9723 | 157.929 | 4.827% | 9.655% | 14.48% | 131.0531 | 144.8482
Zr, ppm | 252.1642 | 11.5551 | 229.0541 | 275.2744 | 217.499 | 286.8295 | 4.582% | 9.165% | 13.75% | 239.556 | 264.7725
Borate / Peroxide Fusion ICP
Al, wt.% | 7.5948 | 0.17451 | 7.2458 | 7.9438 | 7.0712 | 8.1183 | 2.298% | 4.596% | 6.893% | 7.215 | 7.9745
As, ppm | 36.1983 | 4.6495 | 26.8994 | 45.4973 | 22.2499 | 50.1468 | 12.84% | 25.69% | 38.53% | 34.3884 | 38.0082
Ba, ppm | 2795.3371 | 105.4601 | 2584.4169 | 3006.2573 | 2478.9568 | 3111.7174 | 3.773% | 7.545% | 11.32% | 2655.5703 | 2935.104
Be, ppm | 3.0828 | 0.52274 | 2.0373 | 4.1282 | 1.5146 | 4.651 | 16.96% | 33.91% | 50.87% | 2.9286 | 3.2369
Bi, ppm | 6.1971 | 0.35724 | 5.4827 | 6.9116 | 5.1254 | 7.2689 | 5.765% | 11.53% | 17.29% | 5.8873 | 6.507
Ca, wt.% | 0.60822 | 0.028674 | 0.55088 | 0.66557 | 0.5222 | 0.69424 | 4.714% | 9.429% | 14.14% | 0.57781 | 0.63863
Ce, ppm | 95.8806 | 5.9077 | 84.0652 | 107.696 | 78.1575 | 113.6037 | 6.162% | 12.32% | 18.48% | 91.0866 | 100.6746
Co, ppm | 15.2807 | 2.1279 | 11.0248 | 19.5365 | 8.8969 | 21.6644 | 13.93% | 27.85% | 41.78% | 14.5166 | 16.0447
Cs, ppm | 7.0989 | 0.43916 | 6.2206 | 7.9772 | 5.7814 | 8.4164 | 6.186% | 12.37% | 18.56% | 6.7439 | 7.4538
Cu, ppm | 1571.4521 | 52.711 | 1466.0301 | 1676.8741 | 1413.3191 | 1729.5851 | 3.354% | 6.709% | 10.06% | 1492.8795 | 1650.0247
Dy, ppm | 3.7842 | 0.34019 | 3.1038 | 4.4646 | 2.7636 | 4.8048 | 8.99% | 17.98% | 26.97% | 3.595 | 3.9734
Er, ppm | 1.1377 | 0.11657 | 0.90455 | 1.3708 | 0.78798 | 1.4874 | 10.25% | 20.49% | 30.74% | 1.0808 | 1.1946
Eu, ppm | 1.4579 | 0.088585 | 1.2807 | 1.635 | 1.1921 | 1.7236 | 6.076% | 12.15% | 18.23% | 1.385 | 1.5307
Fe, wt.% | 4.2268 | 0.094096 | 4.0386 | 4.415 | 3.9445 | 4.5091 | 2.226% | 4.452% | 6.679% | 4.0155 | 4.4381
Ga, ppm | 25.8159 | 1.4196 | 22.9768 | 28.6551 | 21.5572 | 30.0746 | 5.499% | 11% | 16.5% | 24.5251 | 27.1067
Gd, ppm | 6.4258 | 0.74536 | 4.9351 | 7.9165 | 4.1897 | 8.6619 | 11.6% | 23.2% | 34.8% | 6.1045 | 6.7471
Hf, ppm | 7.2583 | 0.7373 | 5.7837 | 8.7329 | 5.0464 | 9.4702 | 10.16% | 20.32% | 30.47% | 6.8954 | 7.6212
Ho, ppm | 0.53431 | 0.070188 | 0.39394 | 0.67469 | 0.32375 | 0.74488 | 13.14% | 26.27% | 39.41% | 0.5076 | 0.56103
In, ppm | 0.62198 | 0.076009 | 0.46996 | 0.774 | 0.39396 | 0.85001 | 12.22% | 24.44% | 36.66% | 0.59088 | 0.65308
K, wt.% | 2.9394 | 0.12246 | 2.6945 | 3.1844 | 2.5721 | 3.3068 | 4.166% | 8.332% | 12.5% | 2.7925 | 3.0864
La, ppm | 47.9702 | 2.7169 | 42.5365 | 53.4039 | 39.8196 | 56.1208 | 5.664% | 11.33% | 16.99% | 45.5717 | 50.3687
Li, ppm | 20.6086 | 3.1081 | 14.3924 | 26.8247 | 11.2844 | 29.9327 | 15.08% | 30.16% | 45.24% | 19.5781 | 21.639
Mg, wt.% | 0.28787 | 0.015173 | 0.25752 | 0.31821 | 0.24235 | 0.33339 | 5.271% | 10.54% | 15.81% | 0.27348 | 0.30226
Mn, wt.% | 0.039483 | 0.001253 | 0.036977 | 0.041989 | 0.035724 | 0.043242 | 3.174% | 6.347% | 9.521% | 0.037509 | 0.041457
Nb, ppm | 18.5042 | 2.7194 | 13.0655 | 23.943 | 10.3461 | 26.6623 | 14.7% | 29.39% | 44.09% | 17.579 | 19.4295
Nd, ppm | 40.4846 | 1.8185 | 36.8476 | 44.1217 | 35.0291 | 45.9402 | 4.492% | 8.984% | 13.48% | 38.4604 | 42.5089
P, wt.% | 0.02829 | 0.002871 | 0.022548 | 0.034032 | 0.019677 | 0.036903 | 10.15% | 20.3% | 30.45% | 0.026876 | 0.029705
Pb, ppm | 30.8094 | 3.3005 | 24.2083 | 37.4105 | 20.9078 | 40.7111 | 10.71% | 21.43% | 32.14% | 29.269 | 32.3499
Pr, ppm | 10.8876 | 0.73128 | 9.4251 | 12.3502 | 8.6938 | 13.0815 | 6.717% | 13.43% | 20.15% | 10.3433 | 11.432
Rb, ppm | 137.4463 | 4.5166 | 128.4131 | 146.4795 | 123.8965 | 150.9961 | 3.286% | 6.572% | 9.858% | 130.574 | 144.3186
S, wt.% | 0.071317 | 0.013132 | 0.045052 | 0.097581 | 0.03192 | 0.11071 | 18.41% | 36.83% | 55.24% | 0.067751 | 0.074883
Sb, ppm | 1.9566 | 0.32233 | 1.312 | 2.6013 | 0.98963 | 2.9236 | 16.47% | 32.95% | 49.42% | 1.8588 | 2.0545
Si, wt.% | 32.3295 | 1.1333 | 30.0629 | 34.5961 | 28.9296 | 35.7293 | 3.505% | 7.011% | 10.52% | 30.713 | 33.9459
Sm, ppm | 7.6355 | 0.28611 | 7.0633 | 8.2078 | 6.7772 | 8.4939 | 3.747% | 7.494% | 11.24% | 7.2538 | 8.0173
Sn, ppm | 7.4879 | 1.1858 | 5.1163 | 9.8594 | 3.9305 | 11.0452 | 15.84% | 31.67% | 47.51% | 7.1135 | 7.8623
Sr, ppm | 158.6086 | 6.4929 | 145.6227 | 171.5944 | 139.1298 | 178.0874 | 4.094% | 8.187% | 12.28% | 150.6781 | 166.539
Ta, ppm | 1.376 | 0.14387 | 1.0882 | 1.6637 | 0.94434 | 1.8076 | 10.46% | 20.91% | 31.37% | 1.3072 | 1.4447
Tb, ppm | 0.81381 | 0.10046 | 0.61289 | 1.0147 | 0.51243 | 1.1152 | 12.34% | 24.69% | 37.03% | 0.77312 | 0.85451
Th, ppm | 14.6884 | 0.64055 | 13.4073 | 15.9695 | 12.7668 | 16.6101 | 4.361% | 8.722% | 13.08% | 13.954 | 15.4228
Ti, wt.% | 0.12759 | 0.0050742 | 0.11744 | 0.13773 | 0.11236 | 0.14281 | 3.977% | 7.954% | 11.93% | 0.12121 | 0.13396
Tl, ppm | 0.7437 | 0.04362 | 0.65646 | 0.83094 | 0.61284 | 0.87456 | 5.865% | 11.73% | 17.6% | 0.70652 | 0.78089
Tm, ppm | 0.14077 | 0.025741 | 0.089285 | 0.19225 | 0.063544 | 0.21799 | 18.29% | 36.57% | 54.86% | 0.13373 | 0.1478
U, ppm | 5.0041 | 0.4113 | 4.1815 | 5.8267 | 3.7702 | 6.238 | 8.219% | 16.44% | 24.66% | 4.7539 | 5.2543
V, ppm | 11.2958 | 1.5717 | 8.1525 | 14.4391 | 6.5809 | 16.0108 | 13.91% | 27.83% | 41.74% | 10.731 | 11.8606
W, ppm | 3.0182 | 0.45623 | 2.1058 | 3.9307 | 1.6495 | 4.3869 | 15.12% | 30.23% | 45.35% | 2.8673 | 3.1691
Y, ppm | 16.5676 | 1.4194 | 13.7288 | 19.4064 | 12.3094 | 20.8259 | 8.567% | 17.13% | 25.7% | 15.7392 | 17.396
Yb, ppm | 0.75552 | 0.072759 | 0.61 | 0.90104 | 0.53724 | 0.9738 | 9.63% | 19.26% | 28.89% | 0.71774 | 0.7933
Zn, ppm | 138.9064 | 9.9651 | 118.9763 | 158.8366 | 109.0112 | 168.8016 | 7.174% | 14.35% | 21.52% | 131.9611 | 145.8518
Zr, ppm | 270.0568 | 18.6084 | 232.8399 | 307.2737 | 214.2315 | 325.8822 | 6.891% | 13.78% | 20.67% | 256.554 | 283.5597
Aqua Regia Digestion
Ag, ppm | 0.51648 | 0.048596 | 0.41929 | 0.61367 | 0.37069 | 0.66227 | 9.409% | 18.82% | 28.23% | 0.49066 | 0.5423
Al, wt.% | 0.81726 | 0.049255 | 0.71875 | 0.91577 | 0.66949 | 0.96502 | 6.027% | 12.05% | 18.08% | 0.7764 | 0.85812
As, ppm | 31.6861 | 1.8173 | 28.0515 | 35.3206 | 26.2342 | 37.1379 | 5.735% | 11.47% | 17.21% | 30.1018 | 33.2704
Au, ppm | 0.39498 | 0.018885 | 0.35721 | 0.43275 | 0.33833 | 0.45164 | 4.781% | 9.563% | 14.34% | 0.37523 | 0.41473
Ba, ppm | 248.984 | 15.4793 | 218.0254 | 279.9426 | 202.5461 | 295.4219 | 6.217% | 12.43% | 18.65% | 236.5348 | 261.4332
Be, ppm | 0.93372 | 0.079028 | 0.77567 | 1.0918 | 0.69664 | 1.1708 | 8.464% | 16.93% | 25.39% | 0.88704 | 0.98041
Bi, ppm | 5.7447 | 0.35304 | 5.0386 | 6.4508 | 4.6856 | 6.8039 | 6.146% | 12.29% | 18.44% | 5.4575 | 6.032
Ca, wt.% | 0.33678 | 0.017764 | 0.30125 | 0.3723 | 0.28348 | 0.39007 | 5.275% | 10.55% | 15.82% | 0.31994 | 0.35361
Cd, ppm | 0.34212 | 0.02872 | 0.28468 | 0.39956 | 0.25596 | 0.42828 | 8.395% | 16.79% | 25.18% | 0.32501 | 0.35922
Ce, ppm | 79.5606 | 5.2544 | 69.0518 | 90.0693 | 63.7975 | 95.3237 | 6.604% | 13.21% | 19.81% | 75.5825 | 83.5386
Co, ppm | 13.8923 | 0.90601 | 12.0803 | 15.7043 | 11.1743 | 16.6104 | 6.522% | 13.04% | 19.56% | 13.1977 | 14.5869
Cr, ppm | 17.6185 | 1.813 | 13.9925 | 21.2446 | 12.1795 | 23.0576 | 10.29% | 20.58% | 30.87% | 16.7376 | 18.4995
Cs, ppm | 1.3175 | 0.16352 | 0.99048 | 1.6445 | 0.82697 | 1.8081 | 12.41% | 24.82% | 37.23% | 1.2516 | 1.3834
Cu, ppm | 1561.789 | 47.1271 | 1467.5349 | 1656.0432 | 1420.4078 | 1703.1703 | 3.018% | 6.035% | 9.053% | 1483.6996 | 1639.8785
Dy, ppm | 1.7302 | 0.32432 | 1.0815 | 2.3788 | 0.75721 | 2.7031 | 18.75% | 37.49% | 56.24% | 1.6437 | 1.8167
Eu, ppm | 0.91738 | 0.091507 | 0.73436 | 1.1004 | 0.64286 | 1.1919 | 9.975% | 19.95% | 29.92% | 0.87151 | 0.96325
Fe, wt.% | 3.4952 | 0.14925 | 3.1967 | 3.7937 | 3.0474 | 3.9429 | 4.27% | 8.54% | 12.81% | 3.3204 | 3.6699
Ga, ppm | 6.4321 | 0.66058 | 5.1109 | 7.7533 | 4.4504 | 8.4138 | 10.27% | 20.54% | 30.81% | 6.1105 | 6.7537
Gd, ppm | 3.5854 | 0.34381 | 2.8978 | 4.273 | 2.554 | 4.6168 | 9.589% | 19.18% | 28.77% | 3.4061 | 3.7647
Ge, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Hf, ppm | 1.2219 | 0.22161 | 0.77868 | 1.6651 | 0.55707 | 1.8867 | 18.14% | 36.27% | 54.41% | 1.1608 | 1.283
Hg, ppm | < 0.05 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Ho, ppm | 0.22258 | 0.036577 | 0.14943 | 0.29574 | 0.11285 | 0.33232 | 16.43% | 32.87% | 49.3% | 0.21146 | 0.23371
In, ppm | 0.5788 | 0.042398 | 0.49401 | 0.6636 | 0.45161 | 0.706 | 7.325% | 14.65% | 21.98% | 0.54986 | 0.60774
K, wt.% | 0.32201 | 0.023125 | 0.27576 | 0.36826 | 0.25264 | 0.39139 | 7.181% | 14.36% | 21.54% | 0.30591 | 0.33811
La, ppm | 39.8025 | 2.4565 | 34.8894 | 44.7156 | 32.4329 | 47.1721 | 6.172% | 12.34% | 18.52% | 37.8124 | 41.7926
Li, ppm | 4.8746 | 0.66847 | 3.5376 | 6.2115 | 2.8691 | 6.88 | 13.71% | 27.43% | 41.14% | 4.6308 | 5.1183
Lu, ppm | 0.033944 | 0.0053262 | 0.023292 | 0.044596 | 0.017966 | 0.049923 | 15.69% | 31.38% | 47.07% | 0.032247 | 0.035641
Mg, wt.% | 0.15831 | 0.011861 | 0.13459 | 0.18203 | 0.12273 | 0.19389 | 7.492% | 14.98% | 22.48% | 0.1504 | 0.16623
Mn, wt.% | 0.03528 | 0.0011627 | 0.032955 | 0.037606 | 0.031792 | 0.038769 | 3.296% | 6.591% | 9.887% | 0.033516 | 0.037044
Mo, ppm | 3.0239 | 0.17285 | 2.6782 | 3.3696 | 2.5054 | 3.5425 | 5.716% | 11.43% | 17.15% | 2.8727 | 3.1751
Na, wt.% | 0.092062 | 0.0093871 | 0.073288 | 0.11084 | 0.063901 | 0.12022 | 10.2% | 20.39% | 30.59% | 0.087459 | 0.096665
Nd, ppm | 30.4811 | 3.5143 | 23.4524 | 37.5098 | 19.9381 | 41.0241 | 11.53% | 23.06% | 34.59% | 28.957 | 32.0052
Ni, ppm | 8.903 | 0.6374 | 7.6282 | 10.1778 | 6.9908 | 10.8152 | 7.159% | 14.32% | 21.48% | 8.4578 | 9.3481
P, wt.% | 0.023418 | 0.0016927 | 0.020033 | 0.026804 | 0.01834 | 0.028496 | 7.228% | 14.46% | 21.68% | 0.022247 | 0.024589
Pb, ppm | 17.0849 | 1.6681 | 13.7488 | 20.421 | 12.0807 | 22.0891 | 9.763% | 19.53% | 29.29% | 16.2306 | 17.9391
Pr, ppm | 8.4199 | 0.69814 | 7.0236 | 9.8161 | 6.3254 | 10.5143 | 8.292% | 16.58% | 24.87% | 7.9989 | 8.8409
Rb, ppm | 19.2514 | 1.4096 | 16.4322 | 22.0705 | 15.0226 | 23.4801 | 7.322% | 14.64% | 21.97% | 18.2888 | 20.2139
S, wt.% | 0.065101 | 0.0065831 | 0.051935 | 0.078268 | 0.045352 | 0.084851 | 10.11% | 20.22% | 30.34% | 0.061846 | 0.068357
Sb, ppm | 1.1247 | 0.16323 | 0.79819 | 1.4511 | 0.63496 | 1.6143 | 14.51% | 29.03% | 43.54% | 1.0684 | 1.1809
Sc, ppm | 1.8901 | 0.19437 | 1.5013 | 2.2788 | 1.3069 | 2.4732 | 10.28% | 20.57% | 30.85% | 1.7956 | 1.9846
Se, ppm | 2.3169 | 0.43504 | 1.4468 | 3.187 | 1.0118 | 3.622 | 18.78% | 37.55% | 56.33% | 2.2011 | 2.4328
Sm, ppm | 5.0524 | 0.45465 | 4.1431 | 5.9617 | 3.6884 | 6.4163 | 8.999% | 18% | 27% | 4.7998 | 5.305
Sn, ppm | 1.2715 | 0.1375 | 0.99648 | 1.5465 | 0.85898 | 1.684 | 10.81% | 21.63% | 32.44% | 1.2079 | 1.3351
Sr, ppm | 12.3315 | 1.1698 | 9.9918 | 14.6711 | 8.822 | 15.8409 | 9.487% | 18.97% | 28.46% | 11.7149 | 12.948
Tb, ppm | 0.43862 | 0.052166 | 0.33429 | 0.54295 | 0.28212 | 0.59512 | 11.89% | 23.79% | 35.68% | 0.41669 | 0.46055
Te, ppm | 0.065111 | 0.011375 | 0.042361 | 0.087861 | 0.030986 | 0.099236 | 17.47% | 34.94% | 52.41% | 0.061856 | 0.068367
Th, ppm | 8.8816 | 0.70476 | 7.4721 | 10.2912 | 6.7674 | 10.9959 | 7.935% | 15.87% | 23.81% | 8.4376 | 9.3257
Ti, wt.% | 0.019173 | 0.0027197 | 0.013734 | 0.024613 | 0.011014 | 0.027333 | 14.18% | 28.37% | 42.55% | 0.018215 | 0.020132
Tl, ppm | 0.10972 | 0.011326 | 0.087071 | 0.13237 | 0.075745 | 0.1437 | 10.32% | 20.64% | 30.97% | 0.10424 | 0.11521
Tm, ppm | 0.0473 | 0.0078017 | 0.031697 | 0.062904 | 0.023895 | 0.070705 | 16.49% | 32.99% | 49.48% | 0.044935 | 0.049665
U, ppm | 2.3721 | 0.19441 | 1.9832 | 2.7609 | 1.7888 | 2.9553 | 8.196% | 16.39% | 24.59% | 2.2535 | 2.4907
V, ppm | 6.0549 | 0.84668 | 4.3615 | 7.7483 | 3.5149 | 8.595 | 13.98% | 27.97% | 41.95% | 5.7522 | 6.3577
W, ppm | < 1 | IND | IND | IND | IND | IND | IND | IND | IND | IND | IND
Y, ppm | 7.0814 | 0.49223 | 6.0969 | 8.0658 | 5.6047 | 8.5581 | 6.951% | 13.9% | 20.85% | 6.7273 | 7.4354
Yb, ppm | 0.26733 | 0.040945 | 0.18545 | 0.34922 | 0.1445 | 0.39017 | 15.32% | 30.63% | 45.95% | 0.25397 | 0.2807
Zn, ppm | 67.1665 | 5.0341 | 57.0982 | 77.2347 | 52.0641 | 82.2688 | 7.495% | 14.99% | 22.48% | 63.8081 | 70.5248
Zr, ppm | 47.4967 | 4.9543 | 37.5881 | 57.4053 | 32.6338 | 62.3596 | 10.43% | 20.86% | 31.29% | 45.1219 | 49.8715
Sulphuric Acid Leach
Cu-Sol, ppm | 1271.951 | 65.4312 | 1141.0886 | 1402.8135 | 1075.6574 | 1468.2447 | 5.144% | 10.29% | 15.43% | 1208.3535 | 1335.5486
Gas / Liquid Pycnometry
SG, Unity | 2.7473 | 0.04168 | 2.6639 | 2.8306 | 2.6222 | 2.8723 | 1.517% | 3.034% | 4.551% | 2.6099 | 2.8846

## Fire Assay
| Analytical results for Au in OREAS 905 (Certified Value 0.391 ppm) | [33 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | 28 | 29 | 30 | 32 | 33 | 34 | ppm
|  |  | INAA | FA*OES | FA*AAS | FA*OES | FA*MS | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*NAA | FA*MS | FA*AAS | FA*OES | FA*AAS | FA*MS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*OES | FA*AAS | FA*OES | FA*OES | FA*AAS | 3
|  |  | 0.09g | 30g | 30g | 30g | 30g | 30g | 30g | 50g | 40g | 50g | 30g | 30g | 25g | 50g | 30g | 30g | 25g | 20g | 40g | 50g | 30g | 30g | 30g | 30g | 40g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 3
| 1 | 1 | 0.3796 | 0.37 | 0.374 | 0.39 | 0.406 | 0.393 | 0.386 | 0.401 | 0.397 | 0.349 | 0.384 | 0.375 | 0.382 | 0.369 | 0.399 | 0.382 | 0.392 | 0.378 | 0.391 | 0.3791 | 0.404 | 0.383 | 0.439 | 0.394 | 0.389 | 0.392 | 0.387 | 0.395 | 0.39 | 0.364 | 0.394 | 0.397 | 1
| 1 | 2 | 0.4097 | 0.346 | 0.394 | 0.383 | 0.373 | 0.387 | 0.383 | 0.402 | 0.397 | 0.341 | 0.388 | 0.37 | 0.39 | 0.367 | 0.394 | 0.348 | 0.395 | 0.355 | 0.392 | 0.38322 | 0.386 | 0.387 | 0.43 | 0.392 | 0.377 | 0.392 | 0.384 | 0.391 | 0.38 | 0.354 | 0.401 | 0.403 | 0
| 1 | 3 | 0.4203 | 0.361 | 0.328 | 0.381 | 0.401 | 0.39 | 0.397 | 0.4 | 0.403 | 0.274 | 0.401 | 0.362 | 0.395 | 0.376 | 0.397 | 0.396 | 0.384 | 0.35 | 0.391 | 0.38418 | 0.385 | 0.377 | 0.439 | 0.388 | 0.376 | 0.396 | 0.391 | 0.403 | 0.39 | 0.365 | 0.4 | 0.404 | 16
| 1 | 4 | 0.4304 | 0.353 | 0.384 | 0.39 | 0.407 | 0.39 | 0.392 | 0.339 | 0.401 | 0.355 | 0.399 | 0.376 | 0.405 | 0.363 | 0.397 | 0.374 | 0.398 | 0.373 | 0.394 | 0.38388 | 0.38 | 0.371 | 0.421 | 0.4 | 0.396 | 0.388 | 0.392 | 0.396 | 0.4 | 0.37 | 0.398 | 0.406 | 0.39081
| 1 | 5 | 0.3998 | 0.354 | 0.384 | 0.39 | 0.411 | 0.393 | 0.384 | 0.4 | 0.391 | 0.348 | 0.398 | 0.375 | 0.394 | 0.303 | 0.394 | 0.403 | 0.383 | 0.362 | 0.393 | 0.38698 | 0.391 | 0.384 | 0.425 | 0.387 | 0.388 | 0.396 | 0.396 | 0.401 | 0.41 | 0.365 | 0.403 | 0.405 | 7
| 1 | 6 | 0.4171 | 0.343 | 0.38 | 0.396 | 0.394 | 0.394 | 0.381 | 0.395 | 0.39 | 0.342 | 0.403 | 0.381 | 0.384 | 0.375 | 0.391 | 0.385 | 0.391 | 0.358 | 0.392 | 0.39536 | 0.389 | 0.38 | 0.399 | 0.384 | 0.391 | 0.395 | 0.394 | 0.4 | 0.41 | 0.37 | 0.4 | 0.404
|  | 7 | 00: 0.4235
|  | 8 | 00: 0.3908
|  | 9 | 00: 0.4178
|  | 10 | 00: 0.4226
|  | 11 | 00: 0.3792
|  | 12 | 00: 0.3856
|  | 13 | 00: 0.3941
|  | 14 | 00: 0.4024
|  | 15 | 00: 0.4128
|  | 16 | 00: 0.371
|  | 17 | 00: 0.3795
|  | 18 | 00: 0.3988
|  | 19 | 00: 0.4162
|  | 20 | 00: 0.3901
| Mean |  | 0.40207 | 0.3545 | 0.374 | 0.38833 | 0.39867 | 0.39117 | 0.38717 | 0.3895 | 0.3965 | 0.33483 | 0.3955 | 0.37317 | 0.39167 | 0.35883 | 0.39533 | 0.38133 | 0.3905 | 0.36267 | 0.39217 | 0.38545 | 0.38917 | 0.38033 | 0.4255 | 0.39083 | 0.38617 | 0.39317 | 0.39067 | 0.39767 | 0.39667 | 0.36467 | 0.39933 | 0.40317
| Median |  | 0.4011 | 0.3535 | 0.382 | 0.39 | 0.4035 | 0.3915 | 0.385 | 0.4 | 0.397 | 0.345 | 0.3985 | 0.375 | 0.392 | 0.368 | 0.3955 | 0.3835 | 0.3915 | 0.36 | 0.392 | 0.38403 | 0.3875 | 0.3815 | 0.4275 | 0.39 | 0.3885 | 0.3935 | 0.3915 | 0.398 | 0.395 | 0.365 | 0.4 | 0.404
| Std Dev. |  | 0.017749 | 0.0098944 | 0.023461 | 0.005465 | 0.013866 | 0.0026394 | 0.0061128 | 0.024858 | 0.0052058 | 0.030235 | 0.0076616 | 0.0064936 | 0.0083586 | 0.027788 | 0.0028752 | 0.019305 | 0.0059582 | 0.010801 | 0.001169 | 0.005477 | 0.0081833 | 0.0057155 | 0.014883 | 0.0057417 | 0.0079854 | 0.0031252 | 0.0044572 | 0.0044572 | 0.012111 | 0.0058538 | 0.0030768 | 0.0031885
| Rel.Std.Dev. |  | 4.415% | 2.791% | 6.273% | 1.407% | 3.478% | 0.6748% | 1.579% | 6.382% | 1.313% | 9.03% | 1.937% | 1.74% | 2.134% | 7.744% | 0.7273% | 5.062% | 1.526% | 2.978% | 0.2981% | 1.421% | 2.103% | 1.503% | 3.498% | 1.469% | 2.068% | 0.7949% | 1.141% | 1.121% | 3.053% | 1.605% | 0.7705% | 0.7909%
| PDM3 |  | 2.881% | -9.29% | -4.3% | -0.6327% | 2.011% | 0.09232% | -0.9312% | -0.3342% | 1.457% | -14.32% | 1.201% | -4.514% | 0.2203% | -8.181% | 1.158% | -2.424% | -0.07827% | -7.2% | 0.3482% | -1.37% | -0.4194% | -2.68% | 8.878% | 0.007025% | -1.187% | 0.6041% | -0.03562% | 1.756% | 1.5% | -6.689% | 2.182% | 3.163%
| Z-Score (Absolute) |  | NA | 4.18 | 0.86 | 0.27 | 1.52 | 0.06 | 0.4 | 1.03 | 0.67 | 5.05 | 0.56 | 1.76 | 0.12 | 2.39 | 0.54 | 0.31 | 0.02 | 3.24 | 0.17 | 0.6 | 0.17 | 1.19 | 4.03 | 0.02 | 0.52 | 0.29 | 0 | 0.81 | 0.69 | 3.01 | 1 | 1.44
| Analytical results for Pd in OREAS 905 (Indicative Value < 1 ppb) | [33 empty] | Indicative
Pd | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 17; 01: 26 | ppb
|  |  | 00: FA*MS; 01: FA*OES | 0
|  |  | 00: 25g; 01: 40g | 0
| 1 | 1 | 00: < 1; 01: < 5 | 1
| 1 | 2 | 00: < 1; 01: < 5 | 1
| 1 | 3 | 00: < 1; 01: < 5 | 16
| 1 | 4 | 00: < 1; 01: < 5 | < 1
| 1 | 5 | 00: < 1; 01: < 5 | 7
| 1 | 6 | 00: < 1; 01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00 to 01
| Analytical results for Pt in OREAS 905 (Indicative Value < 1 ppb) | [33 empty] | Indicative
Pt | Round | Replicate | 00: Lab; 01: Lab | 1
| No | No | 00: 17; 01: 26 | ppb
|  |  | 00: FA*MS; 01: FA*OES | 0
|  |  | 00: 25g; 01: 40g | 0
| 1 | 1 | 00: < 1; 01: < 5 | 1
| 1 | 2 | 00: < 1; 01: < 5 | 1
| 1 | 3 | 00: < 1; 01: < 5 | 16
| 1 | 4 | 00: < 1; 01: < 5 | < 1
| 1 | 5 | 00: < 1; 01: < 5 | 7
| 1 | 6 | 00: < 1; 01: < 5
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 00 to 01
Note on C31: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 905 (Certified Value 0.518 ppm) | [31 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.48 | 0.54796 | 0.53 | 0.4 | 0.53 | 0.53 | 0.6 | 0.6 | 0.3 | 0.8 | 0.5 | 0.6 | < 3 | 0.33 | 1 | 0.56 | 0.62 | 0.56089 | < 0.5 | 0.8 | 0.7 | 0.5 | 0.54 | 0.51 | < 0.5 | 0.52 | 0.35 | 0.534 |  |  | 1
| 1 | 2 | 0.49 | 0.5508 | 0.49 | 0.5 | 0.53 | 0.58 | 0.6 | 0.8 | 0.4 | 0.9 | 0.47 | 0.4 | < 3 | 0.35 | 1 | 0.57 | 0.61 | 0.61596 | < 0.5 | 1.6 | 0.9 | 0.5 | 0.47 | 0.46 | < 0.5 | 0.56 | 0.37 | 0.627 |  |  | 6
| 1 | 3 | 0.48 | 0.53964 | 0.53 | 0.4 | 0.5 | 0.53 | 0.5 | 0.6 | < 0.3 | 0.8 | 0.48 | 0.5 | < 3 | 0.37 | 1.1 | 0.57 | 0.6 | 0.53058 | < 0.5 | 1.4 | 0.8 | 0.5 | 0.5 | 0.51 | < 0.5 | 0.57 | 0.41 | 0.598 |  |  | 16
| 1 | 4 | 0.47 | 0.52605 | 0.55 | 0.5 | 0.61 | 0.53 | 0.5 | 0.6 | < 0.3 | 0.8 | 0.44 | 0.4 | < 3 | 0.38 | 1.1 | 0.56 | 0.59 | 0.66849 | < 0.5 | 1.2 | 0.7 | 0.5 | 0.51 | 0.49 | < 0.5 | 0.59 | 0.4 | 0.617 |  |  | 0.51764
| 1 | 5 | 0.46 | 0.51689 | 0.51 | 0.5 | 0.51 | 0.53 | 0.6 | 0.8 | < 0.3 | 0.8 | 0.45 | 0.3 | < 3 | 0.4 | 1.1 | 0.53 | 0.59 | 0.57874 | < 0.5 | 1.7 | 0.6 | 0.5 | 0.54 | 0.49 | < 0.5 | 0.53 | 0.39 | 0.608 |  |  | 9
| 1 | 6 | 0.49 | 0.55053 | 0.5 | 0.4 | 0.51 | 0.53 | 0.5 | 0.8 | < 0.3 | 0.9 | 0.5 | 0.4 | < 3 | 0.35 | 1.3 | 0.57 | 0.46 | 0.51781 | < 0.5 | 1.1 | 0.6 | 1 | 0.49 | 0.48 | < 0.5 | 0.59 | 0.4 | 0.625
| Mean |  | 0.47833 | 0.53864 | 0.51833 | 0.45 | 0.53167 | 0.53833 | 0.55 | 0.7 | 0.35 | 0.83333 | 0.47333 | 0.43333 |  | 0.36333 | 1.1 | 0.56 | 0.57833 | 0.57874 |  | 1.3 | 0.71667 | 0.58333 | 0.50833 | 0.49 |  | 0.56 | 0.38667 | 0.6015
| Median |  | 0.48 | 0.5438 | 0.52 | 0.45 | 0.52 | 0.53 | 0.55 | 0.7 | 0.35 | 0.8 | 0.475 | 0.4 |  | 0.36 | 1.1 | 0.565 | 0.595 | 0.56981 |  | 1.3 | 0.7 | 0.5 | 0.505 | 0.49 |  | 0.565 | 0.395 | 0.6125
| Std Dev. |  | 0.01169 | 0.014202 | 0.022286 | 0.054772 | 0.040208 | 0.020412 | 0.054772 | 0.10954 | 0.070711 | 0.05164 | 0.025033 | 0.10328 |  | 0.025033 | 0.10954 | 0.015492 | 0.059133 | 0.056198 |  | 0.33466 | 0.1169 | 0.20412 | 0.027869 | 0.018974 |  | 0.029665 | 0.022509 | 0.034795
| Rel.Std.Dev. |  | 2.444% | 2.637% | 4.3% | 12.17% | 7.563% | 3.792% | 9.959% | 15.65% | 20.2% | 6.197% | 5.289% | 23.83% |  | 6.89% | 9.959% | 2.766% | 10.22% | 9.71% |  | 25.74% | 16.31% | 34.99% | 5.482% | 3.872% |  | 5.297% | 5.821% | 5.785%
| PDM3 |  | -7.593% | 4.058% | 0.1342% | -13.07% | 2.71% | 3.998% | 6.252% | 35.23% | -32.39% | 60.99% | -8.559% | -16.29% |  | -29.81% | 112.5% | 8.184% | 11.73% | 11.8% |  | 151.1% | 38.45% | 12.69% | -1.798% | -5.339% |  | 8.184% | -25.3% | 16.2%
| Z-Score (Absolute) |  | 0.61 | 0.1 | 0.14 | 0.95 | 0.17 | 0 | 0.24 | 2.02 | 2.14 | 3.61 | 0.67 | 1.15 | NA | 1.98 | 6.78 | 0.43 | 0.86 | 0.58 | NA | 9.16 | 1.78 | 0.36 | 0.26 | 0.48 | NA | 0.36 | 1.71 | 1.01
| Analytical results for Al in OREAS 905 (Certified Value 7.42 wt.%) | [31 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 7.05 | 7.6453 | 7.76 | 7.67 | 7.41 | 7.54 | 7.51 | 7.45 | 7.33 | 7.36 | 7.28 | 6.1482 | 5.07 | 7.05 | 5.69 | 7.1405 | 6.22 | 7.284 | 7.8693 | 7.57 | 7.16 | 7.15 | 7.38 | 7.31 | 7.33 | 7.5 | 7.23 | 7.46 |  |  | 1
| 1 | 2 | 7.16 | 7.7741 | 7.98 | 7.71 | 7.37 | 7.52 | 7.64 | 7.48 | 7.33 | 7.95 | 7.03 | 5.452 | 6.03 | 6.9 | 6.13 | 7.1795 | 6.66 | 7.095 | 7.8977 | 7.66 | 7.64 | 7.53 | 7.39 | 7.33 | 7.37 | 7.34 | 7.11 | 7.41 |  |  | 0
| 1 | 3 | 7.74 | 7.772 | 8.04 | 7.48 | 7.25 | 7.6 | 7.63 | 7.6 | 7.27 | 7.51 | 7.3 | 7.5551 | 5.16 | 7.06 | 5.5 | 7.0764 | 5.19 | 6.956 | 7.7601 | 7.5 | 7.13 | 7.22 | 7.35 | 7.51 | 7.48 | 7.55 | 7.22 | 7.52 |  |  | 16
| 1 | 4 | 7.55 | 7.9373 | 7.76 | 7.65 | 7.34 | 7.53 | 7.61 | 7.41 | 7.03 | 7.43 | 7.44 | 7.5198 | 5.84 | 7 | 4.67 | 7.0721 | 6.01 | 6.962 | 7.8393 | 7.53 | 7.02 | 7.41 | 7.26 | 7.34 | 7.55 | 7.56 | 7.27 | 7.62 |  |  | 7.4245
| 1 | 5 | 7.47 | 7.8122 | 8.09 | 7.55 | 7.33 | 7.69 | 7.73 | 7.51 | 6.99 | 7.62 | 6.73 | 7.1105 | 5.85 | 7.02 | 5.65 | 7.1651 | 6.48 | 7.047 | 7.7085 | 7.48 | 7.52 | 7.13 | 7.39 | 7.26 | 7.36 | 7.54 | 7.07 | 7.63 |  |  | 10
| 1 | 6 | 7.51 | 7.7422 | 7.64 | 7.78 | 7.45 | 7.63 | 7.57 | 7.48 | 7.48 | 7.8 | 7.14 | 7.7308 | 6 | 7.07 | 5.38 | 7.1972 | 7.17 | 7.056 | 7.8087 | 7.44 | 6.98 | 7.23 | 7.31 | 7.21 | 7.31 | 7.73 | 7.06 | 7.45
| Mean |  | 7.4133 | 7.7805 | 7.8783 | 7.64 | 7.3583 | 7.585 | 7.615 | 7.4883 | 7.2383 | 7.6117 | 7.1533 | 6.9194 | 5.6583 | 7.0167 | 5.5033 | 7.1385 | 6.2883 | 7.0667 | 7.814 | 7.53 | 7.2417 | 7.2783 | 7.3467 | 7.3267 | 7.4 | 7.5367 | 7.16 | 7.515
| Median |  | 7.49 | 7.773 | 7.87 | 7.66 | 7.355 | 7.57 | 7.62 | 7.48 | 7.3 | 7.565 | 7.21 | 7.3152 | 5.845 | 7.035 | 5.575 | 7.1528 | 6.35 | 7.0515 | 7.824 | 7.515 | 7.145 | 7.225 | 7.365 | 7.32 | 7.365 | 7.545 | 7.165 | 7.49
| Std Dev. |  | 0.25851 | 0.095351 | 0.18225 | 0.10881 | 0.069402 | 0.067157 | 0.073689 | 0.064317 | 0.19041 | 0.22658 | 0.25073 | 0.91733 | 0.42874 | 0.062823 | 0.4814 | 0.05311 | 0.66943 | 0.11978 | 0.070371 | 0.07746 | 0.27309 | 0.15804 | 0.052409 | 0.10211 | 0.094234 | 0.12501 | 0.090774 | 0.09225
| Rel.Std.Dev. |  | 3.487% | 1.226% | 2.313% | 1.424% | 0.9432% | 0.8854% | 0.9677% | 0.8589% | 2.631% | 2.977% | 3.505% | 13.26% | 7.577% | 0.8953% | 8.747% | 0.744% | 10.65% | 1.695% | 0.9006% | 1.029% | 3.771% | 2.171% | 0.7134% | 1.394% | 1.273% | 1.659% | 1.268% | 1.228%
| PDM3 |  | -0.1501% | 4.795% | 6.113% | 2.903% | -0.8909% | 2.162% | 2.566% | 0.86% | -2.507% | 2.521% | -3.652% | -6.803% | -23.79% | -5.493% | -25.88% | -3.852% | -15.3% | -4.819% | 5.246% | 1.421% | -2.462% | -1.968% | -1.048% | -1.317% | -0.3297% | 1.511% | -3.562% | 1.219%
| Z-Score (Absolute) |  | 0.11 | 1.27 | 1.58 | 0.83 | 0.07 | 0.65 | 0.75 | 0.35 | 0.45 | 0.74 | 0.72 | 0.32 | 5.46 | 1.15 | 5.95 | 0.76 | 3.46 | 0.99 | 1.38 | 0.48 | 0.44 | 0.32 | 0.1 | 0.17 | 0.07 | 0.5 | 0.7 | 0.43
| Analytical results for As in OREAS 905 (Certified Value 34.7 ppm) | [31 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 33.4 | 35.1056 | 34 | 38 | 34 | 34.9 | 35 | 36 | 34 | 36 | 34.1 | 28 | 46 | 32 | 24 | 34.9 | 25 | 35 | 35.7001 | 39 | 36.3 | 35 | 28 | 33.4 | 32.5 | 27 | 31.8 | 33 | 34.2 |  | 1
| 1 | 2 | 33.3 | 36.0627 | 35 | 38 | 33.8 | 34.9 | 36 | 39 | 34 | 38 | 33.4 | 31 | 42 | 36 | 19 | 35.7 | 26 | 34.5 | 35.1484 | 39 | 37.5 | 37 | 32 | 30.5 | 34 | 27 | 31.7 | 34 | 34.8 |  | 0
| 1 | 3 | 34.4 | 35.786 | 34 | 35 | 33 | 35 | 34 | 36 | 34 | 36 | 33.8 | 31 | 32 | 36 | 22 | 34.8 | 27 | 34.5 | 34.5708 | 36 | 37.1 | 34 | 34 | 33.8 | 33.3 | 27 | 32.4 | 36 | 34.9 |  | 16
| 1 | 4 | 33.2 | 35.672 | 35 | 36 | 34 | 34.9 | 34 | 37 | 33 | 35 | 32.1 | 31 | 46 | 36 | 21 | 33.8 | 30 | 35 | 35.5632 | 37 | 36.5 | 36 | 35 | 34.7 | 33.7 | 27 | 31.7 | 36 | 34.1 |  | 34.6906
| 1 | 5 | 32.5 | 35.0004 | 33 | 36 | 33.9 | 35.2 | 34 | 39 | 35 | 35 | 32.7 | 30 | 41 | 36 | 18 | 35.4 | 26 | 37.24 | 33.553 | 36 | 36.7 | 33 | 36 | 33.2 | 33.6 | 28 | 32.3 | 35 | 35.4 |  | 11
| 1 | 6 | 33.5 | 35.6643 | 36 | 39 | 33.6 | 34.4 | 34 | 38 | 33 | 35 | 33.4 | 34 | 43 | 36 | 21 | 35 | 29 | 36 | 33.2447 | 37 | 37.2 | 32 | 36 | 33.1 | 32.5 | 25 | 33.6 | 33 | 34.5
| Mean |  | 33.3833 | 35.5485 | 34.5 | 37 | 33.7167 | 34.8833 | 34.5 | 37.5 | 33.8333 | 35.8333 | 33.25 | 30.8333 | 41.6667 | 35.3333 | 20.8333 | 34.9333 | 27.1667 | 35.3733 | 34.63 | 37.3333 | 36.8833 | 34.5 | 33.5 | 33.1167 | 33.2667 | 26.8333 | 32.25 | 34.5 | 34.65
| Median |  | 33.35 | 35.6682 | 34.5 | 37 | 33.85 | 34.9 | 34 | 37.5 | 34 | 35.5 | 33.4 | 31 | 42.5 | 36 | 21 | 34.95 | 26.5 | 35 | 34.8596 | 37 | 36.9 | 34.5 | 34.5 | 33.3 | 33.45 | 27 | 32.05 | 34.5 | 34.65
| Std Dev. |  | 0.61128 | 0.41134 | 1.0488 | 1.5492 | 0.38166 | 0.26394 | 0.83666 | 1.3784 | 0.75277 | 1.169 | 0.73417 | 1.9408 | 5.164 | 1.633 | 2.137 | 0.65013 | 1.9408 | 1.066 | 1.0361 | 1.3663 | 0.45789 | 1.8708 | 3.0822 | 1.4077 | 0.63456 | 0.98319 | 0.7287 | 1.3784 | 0.48477
| Rel.Std.Dev. |  | 1.831% | 1.157% | 3.04% | 4.187% | 1.132% | 0.7566% | 2.425% | 3.676% | 2.225% | 3.262% | 2.208% | 6.294% | 12.39% | 4.622% | 10.26% | 1.861% | 7.144% | 3.013% | 2.992% | 3.66% | 1.241% | 5.423% | 9.201% | 4.251% | 1.907% | 3.664% | 2.26% | 3.995% | 1.399%
| PDM3 |  | -3.768% | 2.473% | -0.5494% | 6.657% | -2.808% | 0.5556% | -0.5494% | 8.098% | -2.471% | 3.294% | -4.153% | -11.12% | 20.11% | 1.853% | -39.95% | 0.6997% | -21.69% | 1.968% | -0.1747% | 7.618% | 6.321% | -0.5494% | -3.432% | -4.537% | -4.105% | -22.65% | -7.035% | -0.5494% | -0.1171%
| Z-Score (Absolute) |  | 0.67 | 0.63 | 0 | 1.51 | 0.47 | 0.23 | 0 | 1.81 | 0.4 | 0.81 | 0.75 | 1.87 | 5.5 | 0.91 | 8.25 | 0.26 | 4.43 | 0.3 | 0.08 | 1.71 | 1.44 | 0 | 0.06 | 0.52 | 0.74 | 4.63 | 1.52 | 0 | 0.09
| Analytical results for B in OREAS 905 (Indicative Value 4 ppm) | [31 empty] | Indicative
B | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 12 |  | ppm
|  |  | 01: 4A*MS |  | 2
2
| 1 | 1 | 01: 3 |  | 1
| 1 | 2 | 01: 2 |  | 3
| 1 | 3 | 01: 8 |  | 16
| 1 | 4 | 01: 4 |  | 4
| 1 | 5 | 01: 2 |  | 9
| 1 | 6 | 01: 5
| Mean |  | 01: 4
| Median |  | 01: 3.5
| Std Dev. |  | 01: 2.2804
| Rel.Std.Dev. |  | 01: 57.01%
PDM3 | bar: 01
Z-Score (Absolute) | bar: 01
| Analytical results for Ba in OREAS 905 (Certified Value 2699 ppm) | [31 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 2600 | 2848.8202 | 2760 | 2760 | 2820 | 2820 | 2850 | 2590 | 2670 | 2753 | 2560 | 2580 | 3194 | 2530 | 2567 | 2623 | 2576 | 2691 | 2747.437 | 2998 | > 1000 | 2726 | 3020 | 2770 | 2740 | 2430 | 2760 | 2576 | 2506 |  | 1
| 1 | 2 | 2660 | 2858.2687 | 2721 | 2800 | 2800 | 2840 | 2870 | 2620 | 2710 | 2939 | 2510 | 2562 | 3217 | 2600 | 2614 | 2637 | 2633 | 2696 | 2750.107 | 2988 | > 1000 | 2763 | 2950 | 2780 | 2760 | 2440 | 2720 | 2522 | 2668 |  | 0
| 1 | 3 | 2860 | 2864.4178 | 2728 | 2700 | 2750 | 2850 | 2850 | 2640 | 2550 | 2730 | 2570 | 2703 | 3070 | 2530 | 2625 | 2609 | 2615 | 2694 | 2704.9071 | 2813 | > 1000 | 2773 | 2770 | 2740 | 2830 | 2470 | 2780 | 2579 | 2652 |  | 16
| 1 | 4 | 2800 | 2855.2614 | 2719 | 2720 | 2790 | 2830 | 2820 | 2640 | 2520 | 2693 | 2570 | 2649 | 3239 | 2560 | 2587 | 2595 | 2555 | 2697 | 2727.897 | 2922 | > 1000 | 2742 | 2880 | 2720 | 2770 | 2500 | 2790 | 2571 | 2641 |  | 2699.3903
| 1 | 5 | 2770 | 2849.7796 | 2687 | 2700 | 2820 | 2880 | 2840 | 2710 | 2590 | 2721 | 2550 | 2505 | 3316 | 2570 | 2576 | 2650 | 2514 | 2693 | 2685.707 | 2833 | > 1000 | 2747 | 2880 | 2770 | 2740 | 2420 | 2770 | 2517 | 2659 |  | 12
| 1 | 6 | 2770 | 2837.3472 | 2644 | 2810 | 2860 | 2860 | 2840 | 2700 | 2590 | 2727 | 2520 | 2611 | 3240 | 2510 | 2492 | 2654 | 2451 | 2658 | 2738.6771 | 2817 | > 1000 | 2717 | 2870 | 2750 | 2720 | 2410 | 2850 | 2518 | 2664
| Mean |  | 2743.3333 | 2852.3158 | 2709.8333 | 2748.3333 | 2806.6667 | 2846.6667 | 2845 | 2650 | 2605 | 2760.5 | 2546.6667 | 2601.6667 | 3212.6667 | 2550 | 2576.8333 | 2628 | 2557.3333 | 2688.1667 | 2725.7887 | 2895.1667 |  | 2744.6667 | 2895 | 2755 | 2760 | 2445 | 2778.3333 | 2547.1667 | 2631.6667
| Median |  | 2770 | 2852.5205 | 2720 | 2740 | 2810 | 2845 | 2845 | 2640 | 2590 | 2728.5 | 2555 | 2595.5 | 3228 | 2545 | 2581.5 | 2630 | 2565.5 | 2693.5 | 2733.287 | 2877.5 |  | 2744.5 | 2880 | 2760 | 2750 | 2435 | 2775 | 2546.5 | 2655.5
| Std Dev. |  | 95.6382 | 9.3089 | 39.7765 | 49.1596 | 36.697 | 21.6025 | 16.4317 | 46.4758 | 72.0417 | 89.5405 | 25.8199 | 69.258 | 81.0424 | 32.8634 | 47.1144 | 23.3067 | 67.2151 | 14.9321 | 25.5858 | 85.6024 |  | 21.2854 | 84.0833 | 22.5832 | 38.4708 | 33.9116 | 42.6224 | 31.0059 | 62.2982
| Rel.Std.Dev. |  | 3.486% | 0.3264% | 1.468% | 1.789% | 1.307% | 0.7589% | 0.5776% | 1.754% | 2.766% | 3.244% | 1.014% | 2.662% | 2.523% | 1.289% | 1.828% | 0.8869% | 2.628% | 0.5555% | 0.9387% | 2.957% |  | 0.7755% | 2.904% | 0.8197% | 1.394% | 1.387% | 1.534% | 1.217% | 2.367%
| PDM3 |  | 1.628% | 5.665% | 0.3869% | 1.813% | 3.974% | 5.456% | 5.394% | -1.83% | -3.497% | 2.264% | -5.658% | -3.62% | 19.01% | -5.534% | -4.54% | -2.645% | -5.263% | -0.4158% | 0.9779% | 7.253% |  | 1.677% | 7.246% | 2.06% | 2.245% | -9.424% | 2.924% | -5.639% | -2.509%
| Z-Score (Absolute) |  | 0.11 | 0.79 | 0.1 | 0.14 | 0.51 | 0.75 | 0.74 | 0.47 | 0.75 | 0 | 1.11 | 0.77 | 3.21 | 1.09 | 0.92 | 0.6 | 1.04 | 0.23 | 0 | 1.06 | NA | 0.12 | 1.05 | 0.18 | 0.22 | 1.74 | 0.33 | 1.11 | 0.43
| Analytical results for Be in OREAS 905 (Certified Value 3.04 ppm) | [31 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
2
| 1 | 1 | 3.22 | 3.0094 | 3.3 | 3 | 2.8 | 3.25 | 3 | 4 | 2.8 | 2.9 | 3 | 3 | < 3 | 3 | 2.7 | 2.75 | 2 | 3 | 3.47 | 5 | < 1 | 3 | 4 | 2.72 | 3.16 | 2.7 | 2.95 | 2.8 | 4 |  | 1
| 1 | 2 | 3.46 | 3.2247 | 3.4 | 3.1 | 2.91 | 3.31 | 3.1 | 4 | 2.8 | 3 | 3 | 3.1 | 3 | 2.7 | 2.7 | 2.89 | 2 | 3 | 3.5197 | 4 | < 1 | 3 | 3.8 | 2.42 | 3.26 | 2.5 | 2.96 | 2.6 | 3 |  | 0
| 1 | 3 | 3.47 | 3.0383 | 3 | 3 | 2.75 | 3.28 | 3 | 3.5 | 2.8 | 2.8 | 3 | 3.1 | < 3 | 2.7 | 2.7 | 2.6 | 2 | 3 | 3.3847 | 4 | < 1 | 3 | 3.8 | 2.93 | 3.19 | 2.7 | 2.98 | 3 | 3 |  | 16
| 1 | 4 | 3.34 | 3.1394 | 3.2 | 3 | 2.78 | 3.23 | 2.9 | 4 | 2.7 | 2.8 | 3 | 3.1 | < 3 | 2.7 | 2.6 | 2.81 | 2 | 3.1 | 3.4625 | 3 | < 1 | 3 | 3.6 | 2.9 | 3.21 | 2.8 | 3.09 | 2.9 | 3 |  | 3.0438
| 1 | 5 | 3.37 | 3.0602 | 3.2 | 3 | 2.76 | 3.18 | 2.9 | 4 | 2.8 | 2.9 | 3 | 2.7 | 3 | 2.8 | 2.7 | 2.76 | 2 | 3.1 | 3.2821 | 4 | < 1 | 3 | 3.5 | 2.59 | 3.33 | 2.7 | 3.02 | 2.8 | 3 |  | 13
| 1 | 6 | 3.26 | 2.8838 | 3.3 | 3.2 | 2.71 | 3.19 | 3 | 4 | 2.8 | 2.9 | 3 | 3 | < 3 | 2.8 | 2.6 | 2.71 | 2 | 3 | 3.1244 | 3 | < 1 | 3 | 3.5 | 3.05 | 3.19 | 2.5 | 3.15 | 3.1 | 4
| Mean |  | 3.3533 | 3.0593 | 3.2333 | 3.05 | 2.785 | 3.24 | 2.9833 | 3.9167 | 2.7833 | 2.8833 | 3 | 3 | 3 | 2.7833 | 2.6667 | 2.7533 | 2 | 3.0333 | 3.3739 | 3.8333 |  | 3 | 3.7 | 2.7683 | 3.2233 | 2.65 | 3.025 | 2.8667 | 3.3333
| Median |  | 3.355 | 3.0493 | 3.25 | 3 | 2.77 | 3.24 | 3 | 4 | 2.8 | 2.9 | 3 | 3.05 | 3 | 2.75 | 2.7 | 2.755 | 2 | 3 | 3.4236 | 4 |  | 3 | 3.7 | 2.81 | 3.2 | 2.7 | 3 | 2.85 | 3
| Std Dev. |  | 0.10191 | 0.11621 | 0.13663 | 0.083666 | 0.068337 | 0.050596 | 0.075277 | 0.20412 | 0.040825 | 0.075277 | 0 | 0.15492 | 0 | 0.1169 | 0.05164 | 0.097297 | 0 | 0.05164 | 0.14773 | 0.75277 |  | 0 | 0.2 | 0.23592 | 0.06186 | 0.12247 | 0.079687 | 0.17512 | 0.5164
| Rel.Std.Dev. |  | 3.039% | 3.798% | 4.226% | 2.743% | 2.454% | 1.562% | 2.523% | 5.212% | 1.467% | 2.611% | 0% | 5.164% | 0% | 4.2% | 1.936% | 3.534% | 0% | 1.702% | 4.379% | 19.64% |  | 0% | 5.405% | 8.522% | 1.919% | 4.622% | 2.634% | 6.109% | 15.49%
| PDM3 |  | 10.17% | 0.5081% | 6.225% | 0.2023% | -8.504% | 6.444% | -1.988% | 28.68% | -8.559% | -5.273% | -1.44% | -1.44% | -1.44% | -8.559% | -12.39% | -9.544% | -34.29% | -0.3452% | 10.84% | 25.94% |  | -1.44% | 21.56% | -9.051% | 5.897% | -12.94% | -0.619% | -5.821% | 9.511%
| Z-Score (Absolute) |  | 1.01 | 0.14 | 0.65 | 0.11 | 0.75 | 0.67 | 0.09 | 2.67 | 0.68 | 0.38 | 0.04 | 0.14 | 0.04 | 0.8 | 1.02 | 0.77 | 2.99 | 0.06 | 1.07 | 1.73 | NA | 0.04 | 2.03 | 0.72 | 0.62 | 1.07 | 0.04 | 0.43 | 0.95
| Analytical results for Bi in OREAS 905 (Certified Value 5.72 ppm) | [31 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.51 | 5.7872 | 5.83 | 5.5 | 5.61 | 5.74 | 5.88 | 5.6 | 5.42 | < 5 | 5.99 | 5.94 | < 20 | 5.35 | < 1 | 6.02 | 1 | < 5 | 6.291 | 11 | 6.7 | < 5 | 6.5 | 5.87 | 5.15 | 4.5 | 5.87 | 5.53 | 6.63 |  | 1
| 1 | 2 | 5.49 | 5.733 | 5.94 | 5.5 | 5.45 | 5.68 | 5.93 | 5.9 | 5.37 | < 5 | 5.84 | 5.8 | < 20 | 5.55 | < 1 | 6.03 | 1 | < 5 | 6.433 | 9 | 6.6 | < 5 | 5.7 | 5.1 | 5.06 | 4.6 | 5.64 | 5.53 | 6.82 |  | 0
| 1 | 3 | 5.67 | 5.7821 | 5.93 | 5.5 | 5.3 | 5.83 | 5.89 | 5.7 | 5.35 | < 5 | 5.87 | 6.14 | < 20 | 5.42 | < 1 | 5.88 | 1 | < 5 | 6.483 | 7 | 6.5 | < 5 | 5.8 | 6.03 | 5.07 | 4.3 | 5.81 | 5.45 | 6.8 |  | 16
| 1 | 4 | 5.46 | 5.773 | 5.89 | 5.7 | 5.55 | 5.69 | 5.74 | 5.5 | 5.16 | < 5 | 5.77 | 6.03 | < 20 | 5.43 | < 1 | 5.96 | 2 | < 5 | 6.27 | 6 | 6.4 | < 5 | 5.6 | 5.75 | 5.13 | 4.5 | 5.79 | 5.55 | 6.64 |  | 5.7152
| 1 | 5 | 5.67 | 5.691 | 5.84 | 5.7 | 5.55 | 5.8 | 5.9 | 5.7 | 5.3 | 5 | 5.79 | 5.74 | < 20 | 5.49 | < 1 | 6.1 | 1 | < 5 | 6.137 | 7 | 6.5 | < 5 | 6.1 | 5.67 | 5.17 | 4.3 | 5.85 | 5.48 | 6.65 |  | 14
| 1 | 6 | 5.5 | 5.7793 | 5.69 | 5.6 | 5.5 | 5.92 | 5.88 | 5.6 | 5.35 | 5 | 5.85 | 6.33 | < 20 | 5.51 | < 1 | 5.9 | < 1 | < 5 | 6.132 | < 5 | 6.3 | < 5 | 5.8 | 5.7 | 5.12 | 4.3 | 5.99 | 5.37 | 6.74
| Mean |  | 5.55 | 5.7576 | 5.8533 | 5.5833 | 5.4933 | 5.7767 | 5.87 | 5.6667 | 5.325 | 5 | 5.8517 | 5.9967 |  | 5.4583 |  | 5.9817 | 1.2 |  | 6.291 | 8 | 6.5 |  | 5.9167 | 5.6867 | 5.1167 | 4.4167 | 5.825 | 5.485 | 6.7133
| Median |  | 5.505 | 5.7762 | 5.865 | 5.55 | 5.525 | 5.77 | 5.885 | 5.65 | 5.35 | 5 | 5.845 | 5.985 |  | 5.46 |  | 5.99 | 1 |  | 6.2805 | 7 | 6.5 |  | 5.8 | 5.725 | 5.125 | 4.4 | 5.83 | 5.505 | 6.695
| Std Dev. |  | 0.094446 | 0.038029 | 0.091797 | 0.098319 | 0.10893 | 0.091797 | 0.066332 | 0.13663 | 0.08961 | 0 | 0.077567 | 0.21933 |  | 0.072226 |  | 0.084004 | 0.44721 |  | 0.14592 | 2 | 0.14142 |  | 0.33116 | 0.31627 | 0.043665 | 0.13292 | 0.1145 | 0.067454 | 0.084774
| Rel.Std.Dev. |  | 1.702% | 0.6605% | 1.568% | 1.761% | 1.983% | 1.589% | 1.13% | 2.411% | 1.683% | 0% | 1.326% | 3.658% |  | 1.323% |  | 1.404% | 37.27% |  | 2.32% | 25% | 2.176% |  | 5.597% | 5.562% | 0.8534% | 3.009% | 1.966% | 1.23% | 1.263%
| PDM3 |  | -2.89% | 0.742% | 2.417% | -2.307% | -3.882% | 1.076% | 2.709% | -0.849% | -6.827% | -12.51% | 2.388% | 4.925% |  | -4.494% |  | 4.663% | -79% |  | 10.08% | 39.98% | 13.73% |  | 3.525% | -0.499% | -10.47% | -22.72% | 1.921% | -4.028% | 17.46%
| Z-Score (Absolute) |  | 0.67 | 0.06 | 0.23 | 0.58 | 0.84 | 0 | 0.28 | 0.33 | 1.25 | 2.31 | 0.22 | 0.65 | NA | 0.95 | NA | 0.61 | 13.62 | NA | 1.53 | 6.61 | 2.15 | NA | 0.07 | 0.08 | 1.96 | 4.05 | 0.14 | 0.87 | 2.79
| Analytical results for Ca in OREAS 905 (Certified Value 0.59 wt.%) | [31 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.59 | 0.62255 | 0.59 | 0.624 | 0.61 | 0.6 | 0.567 | 0.6 | 0.57 | 0.6 | 0.64 | 0.5295 | 0.42 | 0.58 | 0.5 | 0.5952 | 0.55 | 0.575 | 0.57592 | 0.66 | 0.61 | 0.56 | 0.6 | 0.59 | 0.57 | 0.579 | 0.57 | 0.59 | 0.56 |  | 1
| 1 | 2 | 0.6 | 0.6253 | 0.6 | 0.614 | 0.59 | 0.6 | 0.577 | 0.6 | 0.58 | 0.67 | 0.62 | 0.4906 | 0.49 | 0.57 | 0.54 | 0.5956 | 0.57 | 0.536 | 0.57575 | 0.64 | 0.64 | 0.58 | 0.61 | 0.59 | 0.58 | 0.583 | 0.57 | 0.58 | 0.56 |  | 0
| 1 | 3 | 0.64 | 0.62812 | 0.62 | 0.586 | 0.58 | 0.6 | 0.573 | 0.61 | 0.57 | 0.61 | 0.65 | 0.6085 | 0.44 | 0.58 | 0.47 | 0.5858 | 0.45 | 0.512 | 0.56242 | 0.6 | 0.6 | 0.56 | 0.61 | 0.59 | 0.59 | 0.588 | 0.58 | 0.59 | 0.57 |  | 16
| 1 | 4 | 0.63 | 0.63039 | 0.6 | 0.619 | 0.59 | 0.6 | 0.571 | 0.59 | 0.54 | 0.6 | 0.65 | 0.6039 | 0.48 | 0.57 | 0.42 | 0.5925 | 0.52 | 0.517 | 0.5673 | 0.63 | 0.61 | 0.56 | 0.61 | 0.58 | 0.58 | 0.591 | 0.58 | 0.59 | 0.57 |  | 0.59011
| 1 | 5 | 0.62 | 0.62165 | 0.62 | 0.615 | 0.59 | 0.61 | 0.566 | 0.64 | 0.55 | 0.59 | 0.63 | 0.5707 | 0.5 | 0.58 | 0.53 | 0.5992 | 0.53 | 0.577 | 0.55802 | 0.61 | 0.65 | 0.55 | 0.6 | 0.59 | 0.57 | 0.578 | 0.58 | 0.58 | 0.57 |  | 15
| 1 | 6 | 0.63 | 0.63326 | 0.57 | 0.621 | 0.6 | 0.61 | 0.574 | 0.62 | 0.58 | 0.59 | 0.64 | 0.6125 | 0.5 | 0.56 | 0.48 | 0.5957 | 0.55 | 0.526 | 0.56697 | 0.6 | 0.57 | 0.57 | 0.61 | 0.58 | 0.57 | 0.582 | 0.6 | 0.58 | 0.57
| Mean |  | 0.61833 | 0.62688 | 0.6 | 0.61317 | 0.59333 | 0.60333 | 0.57133 | 0.61 | 0.565 | 0.61 | 0.63833 | 0.56928 | 0.47167 | 0.57333 | 0.49 | 0.594 | 0.52833 | 0.5405 | 0.56773 | 0.62333 | 0.61333 | 0.56333 | 0.60667 | 0.58667 | 0.57667 | 0.5835 | 0.58 | 0.585 | 0.56667
| Median |  | 0.625 | 0.62671 | 0.6 | 0.617 | 0.59 | 0.6 | 0.572 | 0.605 | 0.57 | 0.6 | 0.64 | 0.5873 | 0.485 | 0.575 | 0.49 | 0.5954 | 0.54 | 0.531 | 0.56713 | 0.62 | 0.61 | 0.56 | 0.61 | 0.59 | 0.575 | 0.5825 | 0.58 | 0.585 | 0.57
| Std Dev. |  | 0.019408 | 0.0045414 | 0.018974 | 0.013819 | 0.010328 | 0.005164 | 0.0042269 | 0.017889 | 0.016432 | 0.030332 | 0.01169 | 0.049759 | 0.033714 | 0.008165 | 0.043818 | 0.004548 | 0.042151 | 0.028697 | 0.0071355 | 0.024221 | 0.028752 | 0.010328 | 0.005164 | 0.005164 | 0.008165 | 0.0050892 | 0.010954 | 0.0054772 | 0.005164
| Rel.Std.Dev. |  | 3.139% | 0.7245% | 3.162% | 2.254% | 1.741% | 0.8559% | 0.7398% | 2.933% | 2.908% | 4.972% | 1.831% | 8.741% | 7.148% | 1.424% | 8.942% | 0.7657% | 7.978% | 5.309% | 1.257% | 3.886% | 4.688% | 1.833% | 0.8512% | 0.8802% | 1.416% | 0.8722% | 1.889% | 0.9363% | 0.9113%
| PDM3 |  | 4.783% | 6.231% | 1.676% | 3.908% | 0.5467% | 2.241% | -3.181% | 3.371% | -4.255% | 3.371% | 8.172% | -3.529% | -20.07% | -2.843% | -16.96% | 0.6597% | -10.47% | -8.406% | -3.792% | 5.631% | 3.936% | -4.537% | 2.806% | -0.583% | -2.278% | -1.12% | -1.713% | -0.8655% | -3.972%
| Z-Score (Absolute) |  | 1.22 | 1.54 | 0.55 | 1.23 | 0.31 | 0.67 | 0.5 | 0.92 | 0.74 | 0.48 | 1.96 | 0 | 4.16 | 0.43 | 3.49 | 0.33 | 1.51 | 1.64 | 0.64 | 1.41 | 1.04 | 0.8 | 0.8 | 0.06 | 0.31 | 0.06 | 0.18 | 0 | 0.67
| Analytical results for Cd in OREAS 905 (Certified Value 0.36 ppm) | [31 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 0.32 | 0.33168 | 0.33 | < 1 | 0.3 | 0.33 | 0.33 | < 0.5 | 0.58 | 1 | 0.4 | 0.41 | < 3 | 0.34 | 0.43 | 0.58 | 0.4515 | 2.3 | 1.2 | < 0.4 | < 0.5 | 0.33 | 0.32 | < 1 | 0.33 | 0.39 | 0.42 |  |  |  | 1
| 1 | 2 | 0.34 | 0.30408 | 0.37 | < 1 | 0.34 | 0.34 | 0.35 | < 0.5 | 0.45 | < 1 | 0.4 | 0.38 | < 3 | 0.32 | 0.32 | 0.5 | 0.48467 | 0.8 | 0.9 | < 0.4 | < 0.5 | 0.3 | 0.33 | < 1 | 0.33 | 0.35 | 0.46 |  |  |  | 8
| 1 | 3 | 0.34 | 0.35377 | 0.36 | < 1 | 0.28 | 0.35 | 0.32 | < 0.5 | 0.49 | < 1 | 0.3 | 0.35 | < 3 | 0.34 | 0.37 | 0.66 | 0.36536 | < 0.4 | 0.9 | < 0.4 | < 0.5 | 0.35 | 0.32 | < 1 | 0.32 | 0.36 | 0.54 |  |  |  | 16
| 1 | 4 | 0.32 | 0.37461 | 0.36 | < 1 | 0.33 | 0.28 | 0.39 | < 0.5 | 0.46 | < 1 | 0.4 | 0.35 | < 3 | 0.37 | 0.32 | 0.58 | 0.40416 | < 0.4 | 0.7 | < 0.4 | < 0.5 | 0.34 | 0.32 | < 1 | 0.34 | 0.37 | 0.44 |  |  |  | 0.35727
| 1 | 5 | 0.33 | 0.32576 | 0.34 | < 1 | 0.33 | 0.38 | 0.36 | < 0.5 | 0.53 | 1 | 0.5 | 0.35 | < 3 | 0.34 | 0.39 | 0.66 | 0.44228 | < 0.4 | 0.7 | < 0.4 | < 0.5 | 0.32 | 0.32 | < 1 | 0.34 | 0.35 | 0.46 |  |  |  | 16
| 1 | 6 | 0.31 | 0.35679 | 0.4 | < 1 | 0.35 | 0.35 | 0.37 | < 0.5 | 0.52 | 1 | 0.4 | 0.35 | < 3 | 0.31 | 0.34 | 0.5 | 0.40349 | < 0.4 | 0.7 | < 0.4 | < 0.5 | 0.33 | 0.32 | < 1 | 0.45 | 0.35 | 0.48
| Mean |  | 0.32667 | 0.34111 | 0.36 |  | 0.32167 | 0.33833 | 0.35333 |  | 0.505 | 1 | 0.4 | 0.365 |  | 0.33667 | 0.36167 | 0.58 | 0.42524 | 1.55 | 0.85 |  |  | 0.32833 | 0.32167 |  | 0.35167 | 0.36167 | 0.46667
| Median |  | 0.325 | 0.34272 | 0.36 |  | 0.33 | 0.345 | 0.355 |  | 0.505 | 1 | 0.4 | 0.35 |  | 0.34 | 0.355 | 0.58 | 0.42322 | 1.55 | 0.8 |  |  | 0.33 | 0.32 |  | 0.335 | 0.355 | 0.46
| Std Dev. |  | 0.012111 | 0.025391 | 0.024495 |  | 0.026394 | 0.033116 | 0.02582 |  | 0.048477 | 0 | 0.063246 | 0.0251 |  | 0.020656 | 0.043551 | 0.071554 | 0.042459 | 1.0607 | 0.19748 |  |  | 0.017224 | 0.0040825 |  | 0.048751 | 0.016021 | 0.041312
| Rel.Std.Dev. |  | 3.707% | 7.444% | 6.804% |  | 8.206% | 9.788% | 7.308% |  | 9.599% | 0% | 15.81% | 6.877% |  | 6.135% | 12.04% | 12.34% | 9.985% | 68.43% | 23.23% |  |  | 5.246% | 1.269% |  | 13.86% | 4.43% | 8.853%
| PDM3 |  | -8.566% | -4.523% | 0.7635% |  | -9.966% | -5.301% | -1.103% |  | 41.35% | 179.9% | 11.96% | 2.163% |  | -5.767% | 1.23% | 62.34% | 19.02% | 333.8% | 137.9% |  |  | -8.1% | -9.966% |  | -1.569% | 1.23% | 30.62%
| Z-Score (Absolute) |  | 0.67 | 0.38 | 0 | NA | 0.78 | 0.44 | 0.13 | NA | 2.93 | 12.95 | 0.81 | 0.08 | NA | 0.47 | 0.03 | 4.45 | 1.32 | 24.07 | 9.91 | NA | NA | 0.64 | 0.78 | NA | 0.57 | 0.03 | 1.86
| Analytical results for Ce in OREAS 905 (Certified Value 92 ppm) | [31 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [9 empty] | 0
1
| 1 | 1 | 87.6 | 80.99 | 98.1 | 94 | 93.1 | 85.6 | 91.9 | 91 | 89.1 | 81.5 | 87.1 | 97.54 | 92.2809 | 107 | 92.9 | 97.2 | 84.6 | 90 | 90.4 | 92.8 | 97.6 | [9 empty] | 1
| 1 | 2 | 89.4 | 84.55 | 99 | 93.4 | 93.2 | 85 | 91.8 | 107 | 85.8 | 78.2 | 92.9 | 101.31 | 93.9043 | 125 | 89.8 | 88 | 87.5 | 90 | 90.4 | 90.1 | 97.64 | [9 empty] | 0
| 1 | 3 | 93.5 | 85.41 | 97.9 | 90.5 | 94.4 | 83.7 | 90.4 | 97 | 85.9 | 93.5 | 88.4 | 97.44 | 94.1783 | 110 | 91.5 | 94.6 | 89.7 | 90 | 92.3 | 90.2 | 99.11 | [9 empty] | 16
| 1 | 4 | 90.6 | 79.68 | 101 | 92.6 | 93.6 | 82.5 | 88.3 | 95 | 88.8 | 92.9 | 91.7 | 98.3 | 97.3765 | 111 | 91.2 | 96.1 | 86.2 | 90 | 90.3 | 89.8 | 98.99 | [9 empty] | 92.1763
| 1 | 5 | 90.5 | 77.74 | 98.5 | 93.4 | 95.2 | 90.6 | 91.5 | 92 | 68.7 | 88.2 | 92.8 | 98.19 | 92.2426 | 116 | 89.2 | 96.9 | 87.5 | 90 | 93.5 | 92.5 | 100.57 | [9 empty] | 17
| 1 | 6 | 89.8 | 80.15 | 98.9 | 91.8 | 93.7 | 86 | 92 | 93 | 80.8 | 96.2 | 91.4 | 98.56 | 92.4348 | 92 | 91.2 | 97.6 | 84.5 | 90 | 95 | 89.9 | 100.5
| Mean |  | 90.2333 | 81.42 | 98.9 | 92.6167 | 93.8667 | 85.5667 | 90.9833 | 95.8333 | 83.1833 | 88.4167 | 90.7167 | 98.5567 | 93.7362 | 110.1667 | 90.9667 | 95.0667 | 86.6667 | 90 | 91.9833 | 90.8833 | 99.0683
| Median |  | 90.15 | 80.57 | 98.7 | 93 | 93.65 | 85.3 | 91.65 | 94 | 85.85 | 90.55 | 91.55 | 98.245 | 93.1696 | 110.5 | 91.2 | 96.5 | 86.85 | 90 | 91.35 | 90.15 | 99.05
| Std Dev. |  | 1.9336 | 2.9693 | 1.1153 | 1.2875 | 0.79917 | 2.7833 | 1.4386 | 5.8793 | 7.6971 | 7.1943 | 2.4078 | 1.4186 | 1.9758 | 10.9072 | 1.3125 | 3.6209 | 1.9886 | 0 | 1.9671 | 1.379 | 1.3045
| Rel.Std.Dev. |  | 2.143% | 3.647% | 1.128% | 1.39% | 0.8514% | 3.253% | 1.581% | 6.135% | 9.253% | 8.137% | 2.654% | 1.439% | 2.108% | 9.901% | 1.443% | 3.809% | 2.295% | 0% | 2.139% | 1.517% | 1.317%
| PDM3 |  | -2.108% | -11.67% | 7.294% | 0.4777% | 1.834% | -7.171% | -1.294% | 3.967% | -9.756% | -4.079% | -1.584% | 6.922% | 1.692% | 19.52% | -1.312% | 3.136% | -5.977% | -2.361% | -0.2094% | -1.403% | 7.477%
| Z-Score (Absolute) |  | 0.39 | 3.07 | 2.25 | 0.33 | 0.71 | 1.81 | 0 | 0.63 | 1.66 | 0.32 | 0.24 | 1.97 | 0.67 | 5.67 | 0.17 | 1.51 | 1.48 | 0.46 | 0.14 | 0.19 | 2.3
| Analytical results for Co in OREAS 905 (Certified Value 14.8 ppm) | [31 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 15.2 | 15.2708 | 15.5 | 14 | 14.2 | 15.4 | 16 | 16 | 15.4 | 14 | 14.1 | 12.9 | 14 | 15 | 15 | 14.8 | 16 | 11.5 | 14.869 | 15 | 15.4 | 15 | 15 | 15.1 | 13.4 | 13 | 14 | 15.3 | 15.7 |  | 1
| 1 | 2 | 15.3 | 15.2853 | 15.8 | 15 | 14 | 15.5 | 16 | 16 | 15.1 | 15 | 14.2 | 12.8 | 14 | 14.7 | 15 | 15.2 | 16 | 10.4 | 14.8086 | 15 | 15.9 | 14 | 20 | 13.5 | 13.9 | 14 | 15.1 | 15.2 | 16.7 |  | 0
| 1 | 3 | 15.9 | 15.3547 | 15.6 | 14 | 13.5 | 15.7 | 15 | 16 | 15 | 14 | 14.5 | 13.8 | 14 | 14.6 | 16 | 14.7 | 15 | 11.5 | 14.2206 | 15 | 16.2 | 14 | 20 | 14.8 | 13.7 | 14 | 14.6 | 15.7 | 16.4 |  | 16
| 1 | 4 | 15.1 | 15.6108 | 15.5 | 14 | 14 | 15.2 | 16 | 16 | 14.4 | 14 | 14.4 | 13.3 | 14 | 14.6 | 16 | 14.9 | 15 | 11.5 | 14.5162 | 15 | 15.6 | 14 | 15 | 15.1 | 13.8 | 15 | 14 | 15.5 | 15.8 |  | 14.8446
| 1 | 5 | 15 | 14.9454 | 15.8 | 15 | 14 | 15.4 | 16 | 16 | 15.1 | 14 | 14.4 | 12.2 | 14 | 14.8 | 15 | 15 | 16 | 11.5 | 14.3227 | 14 | 15.7 | 14 | 15 | 14.9 | 14.2 | 15 | 14.3 | 15.3 | 16.2 |  | 18
| 1 | 6 | 15 | 15.787 | 15.8 | 14 | 13.8 | 15.5 | 16 | 16 | 14.9 | 14 | 14.1 | 13.5 | 14 | 14.8 | 15 | 14.5 | 15 | 10.4 | 13.9554 | 14 | 15.8 | 14 | 20 | 15.1 | 13.5 | 14 | 14.6 | 15.1 | 16.1
| Mean |  | 15.25 | 15.3757 | 15.6667 | 14.3333 | 13.9167 | 15.45 | 15.8333 | 16 | 14.9833 | 14.1667 | 14.2833 | 13.0833 | 14 | 14.75 | 15.3333 | 14.85 | 15.5 | 11.1333 | 14.4488 | 14.6667 | 15.7667 | 14.1667 | 17.5 | 14.75 | 13.75 | 14.1667 | 14.4333 | 15.35 | 16.15
| Median |  | 15.15 | 15.32 | 15.7 | 14 | 14 | 15.45 | 16 | 16 | 15.05 | 14 | 14.3 | 13.1 | 14 | 14.75 | 15 | 14.85 | 15.5 | 11.5 | 14.4194 | 15 | 15.75 | 14 | 17.5 | 15 | 13.75 | 14 | 14.45 | 15.3 | 16.15
| Std Dev. |  | 0.33912 | 0.29299 | 0.15055 | 0.5164 | 0.24014 | 0.16432 | 0.40825 | 0 | 0.33116 | 0.40825 | 0.17224 | 0.57067 | 0 | 0.15166 | 0.5164 | 0.2429 | 0.54772 | 0.56804 | 0.35271 | 0.5164 | 0.27325 | 0.40825 | 2.7386 | 0.6253 | 0.2881 | 0.75277 | 0.42269 | 0.21679 | 0.37283
| Rel.Std.Dev. |  | 2.224% | 1.906% | 0.961% | 3.603% | 1.726% | 1.064% | 2.578% | 0% | 2.21% | 2.882% | 1.206% | 4.362% | 0% | 1.028% | 3.368% | 1.636% | 3.534% | 5.102% | 2.441% | 3.521% | 1.733% | 2.882% | 15.65% | 4.239% | 2.095% | 5.314% | 2.929% | 1.412% | 2.309%
| PDM3 |  | 2.731% | 3.578% | 5.538% | -3.444% | -6.251% | 4.078% | 6.66% | 7.783% | 0.9345% | -4.567% | -3.781% | -11.86% | -5.69% | -0.6373% | 3.292% | 0.03632% | 4.415% | -25% | -2.667% | -1.199% | 6.211% | -4.567% | 17.89% | -0.6373% | -7.374% | -4.567% | -2.771% | 3.405% | 8.794%
| Z-Score (Absolute) |  | 0.28 | 0.55 | 0.85 | 0.54 | 0.88 | 0.62 | 1.02 | 1.19 | 0.14 | 0.71 | 0.59 | 1.65 | 0.88 | 0.1 | 0.5 | 0 | 0.67 | 3.86 | 0.42 | 0.19 | 0.95 | 0.71 | 2.75 | 0.16 | 1.14 | 0.71 | 0.43 | 0.52 | 1.35
| Analytical results for Cr in OREAS 905 (Certified Value 19.2 ppm) | [31 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 1
1
| 1 | 1 | 19 | 23.1046 | 21 | 15 | 19 | 19 | 20 | 10 | 13 | 20 | 22 | 18 | 10 | N.A. | 23 | 16 | 31 | 24.4 | 19.8553 | 16 | 24 | 13 | 18 | 21 | 20 | 18 | 20 | 18 |  |  | 1
| 1 | 2 | 18 | 21.3189 | 21 | 15 | 19 | 19 | 20 | 10 | 13 | 20 | 26 | 18 | 10 | N.A. | 18 | 17 | 25 | 26.5 | 19.1425 | 16 | 20 | 16 | 18 | 18 | 10 | 18 | 16 | 18 |  |  | 9
| 1 | 3 | 19 | 21.9495 | 20 | 15 | 19 | 19 | 20 | 10 | 10 | 18 | 20 | 12 | 10 | N.A. | 23 | 19 | 26 | 26.3 | 18.5179 | 17 | 24 | 14 | 21 | 18 | 20 | 18 | 15 | 22 |  |  | 16
| 1 | 4 | 19 | 21.0935 | 21 | 15 | 19 | 19 | 22 | 10 | 13 | 19 | 25 | 8 | 8 | N.A. | 20 | 19 | 27 | 24.5 | 20.7933 | 18 | 21 | 17 | 17 | 18 | 20 | 18 | 14 | 22 |  |  | 19.1784
| 1 | 5 | 20 | 21.0367 | 22 | 15 | 19 | 19 | 22 | 10 | 14 | 20 | 25 | 8 | 10 | N.A. | 20 | 20 | 27 | 24.3 | 18.9538 | 17 | 24 | 14 | 19 | 19 | 20 | 20 | 16 | 18 |  |  | 19
| 1 | 6 | 21 | 23.2891 | 20 | 15 | 19 | 19 | 21 | 10 | 12 | 20 | 24 | 7 | 11 | N.A. | 32 | 18 | 26 | 26.5 | 18.9741 | 17 | 24 | 14 | 25 | 20 | 20 | 19 | 18 | 20
| Mean |  | 19.3333 | 21.9654 | 20.8333 | 15 | 19 | 19 | 20.8333 | 10 | 12.5 | 19.5 | 23.6667 | 11.8333 | 9.8333 |  | 22.6667 | 18.1667 | 27 | 25.4167 | 19.3728 | 16.8333 | 22.8333 | 14.6667 | 19.6667 | 19 | 18.3333 | 18.5 | 16.5 | 19.6667
| Median |  | 19 | 21.6342 | 21 | 15 | 19 | 19 | 20.5 | 10 | 13 | 20 | 24.5 | 10 | 10 |  | 21.5 | 18.5 | 26.5 | 25.4 | 19.0583 | 17 | 24 | 14 | 18.5 | 18.5 | 20 | 18 | 16 | 19
| Std Dev. |  | 1.0328 | 1.009 | 0.75277 | 0 | 0 | 0 | 0.98319 | 0 | 1.3784 | 0.83666 | 2.2509 | 5.0761 | 0.98319 |  | 4.9666 | 1.472 | 2.0976 | 1.1179 | 0.82086 | 0.75277 | 1.8348 | 1.5055 | 2.9439 | 1.2649 | 4.0825 | 0.83666 | 2.1679 | 1.9664
| Rel.Std.Dev. |  | 5.342% | 4.594% | 3.613% | 0% | 0% | 0% | 4.719% | 0% | 11.03% | 4.291% | 9.511% | 42.9% | 9.999% |  | 21.91% | 8.103% | 7.769% | 4.398% | 4.237% | 4.472% | 8.036% | 10.27% | 14.97% | 6.657% | 22.27% | 4.522% | 13.14% | 9.999%
| PDM3 |  | 0.8076% | 14.53% | 8.629% | -21.79% | -0.9304% | -0.9304% | 8.629% | -47.86% | -34.82% | 1.677% | 23.4% | -38.3% | -48.73% |  | 18.19% | -5.276% | 40.78% | 32.53% | 1.014% | -12.23% | 19.06% | -23.53% | 2.546% | -0.9304% | -4.407% | -3.538% | -13.97% | 2.546%
| Z-Score (Absolute) |  | 0.12 | 1.09 | 0.67 | 1.47 | 0 | 0 | 0.67 | 3.31 | 2.21 | 0.18 | 1.72 | 2.64 | 3.37 | NA | 0.66 | 0.31 | 2.65 | 2.36 | 0.03 | 0.8 | 1.62 | 1.59 | 0.15 | 0 | 0.25 | 0.18 | 0.92 | 0.25
| Analytical results for Cs in OREAS 905 (Certified Value 6.78 ppm) | [31 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 17 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 2
3
| 1 | 1 | 6.44 | 6.2097 | 6.75 | 6.38 | 6.73 | 6.57 | 6.2 | 7.6 | 6.9 | 6.09 | 6.5 | 7 | 6.95 | 7.48 | 8.2 | 7.6 | 6.71 | 6.65 | 4.55 | 6.59 | 7 | 6.9 | [8 empty] | 1
| 1 | 2 | 6.49 | 6.1274 | 6.82 | 6.71 | 6.74 | 6.63 | 6.5 | 7.5 | 6.92 | 6.37 | 6.7 | 6 | 7.04 | 7.467 | 8.1 | 7.3 | 6.14 | 6.94 | 4.71 | 6.44 | 7 | 7 | [8 empty] | 0
| 1 | 3 | 6.82 | 6.205 | 6.64 | 6.63 | 6.6 | 6.67 | 6 | 7.5 | 6.8 | 6.28 | 7 | 6 | 7.02 | 7.369 | 7.8 | 7.3 | 6.75 | 7.03 | 4.79 | 6.61 | 7 | 7.1 | [8 empty] | 16
| 1 | 4 | 6.46 | 6.3279 | 6.63 | 6.88 | 6.71 | 6.59 | 6.2 | 7.5 | 6.61 | 6.48 | 7.2 | 6 | 6.94 | 7.13 | 7.6 | 7.4 | 6.95 | 6.83 | 4.84 | 6.53 | 7 | 7 | [8 empty] | 6.7784
| 1 | 5 | 6.64 | 6.2694 | 6.64 | 6.81 | 6.78 | 6.8 | 6.3 | 7.5 | 6.86 | 6.25 | 6.5 | 6 | 7.15 | 7.338 | 7.7 | 7.5 | 6.69 | 6.99 | 4.7 | 6.66 | 7 | 7.1 | [8 empty] | 20
| 1 | 6 | 6.68 | 6.3369 | 6.41 | 6.77 | 6.67 | 6.61 | 6.3 | 7.6 | 6.89 | 6.21 | 7.2 | 6 | 7.11 | 7.358 | 7.6 | 7.4 | 6.75 | 6.68 | 4.69 | 6.89 | 7 | 7.1
| Mean |  | 6.5883 | 6.2461 | 6.6483 | 6.6967 | 6.705 | 6.645 | 6.25 | 7.5333 | 6.83 | 6.28 | 6.85 | 6.1667 | 7.035 | 7.357 | 7.8333 | 7.4167 | 6.665 | 6.8533 | 4.7133 | 6.62 | 7 | 7.0333
| Median |  | 6.565 | 6.2396 | 6.64 | 6.74 | 6.72 | 6.62 | 6.25 | 7.5 | 6.875 | 6.265 | 6.85 | 6 | 7.03 | 7.3635 | 7.75 | 7.4 | 6.73 | 6.885 | 4.705 | 6.6 | 7 | 7.05
| Std Dev. |  | 0.15026 | 0.080723 | 0.13934 | 0.17705 | 0.062849 | 0.083367 | 0.16432 | 0.05164 | 0.11559 | 0.13416 | 0.32711 | 0.40825 | 0.084083 | 0.12591 | 0.2582 | 0.1169 | 0.27348 | 0.16083 | 0.099331 | 0.15232 | 0 | 0.08165
| Rel.Std.Dev. |  | 2.281% | 1.292% | 2.096% | 2.644% | 0.9373% | 1.255% | 2.629% | 0.6855% | 1.692% | 2.136% | 4.775% | 6.62% | 1.195% | 1.711% | 3.296% | 1.576% | 4.103% | 2.347% | 2.107% | 2.301% | 0% | 1.161%
| PDM3 |  | -2.805% | -7.854% | -1.919% | -1.206% | -1.083% | -1.969% | -7.796% | 11.14% | 0.7607% | -7.353% | 1.056% | -9.025% | 3.785% | 8.535% | 15.56% | 9.416% | -1.673% | 1.105% | -30.47% | -2.337% | 3.269% | 3.76%
| Z-Score (Absolute) |  | 0.36 | 1.19 | 0.22 | 0.1 | 0.08 | 0.22 | 1.18 | 1.92 | 0.33 | 1.11 | 0.27 | 1.38 | 0.72 | 1.5 | 2.65 | 1.64 | 0.08 | 0.28 | 4.89 | 0.28 | 0.63 | 0.71
| Analytical results for Cu in OREAS 905 (Certified Value 1533 ppm) | [31 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | 0
0
| 1 | 1 | 1400 | 1540 | 1549.0442 | 1581 | 1440 | 1520 | 1580 | 1550 | 1580 | 1450 | 1610 | 1490 | 1519 | 1610 | 1470 | 1400 | 1474 | 1500 | 1587 | 1573.6041 | 1604 | 1590 | 1560 | 1550 | 1540 | 1490 | 1320 | 1500 | 1460 | 1501.07 | 1
| 1 | 2 | 1440 | 1480 | 1525.3647 | 1548 | 1490 | 1495 | 1570 | 1580 | 1560 | 1430 | 1623 | 1480 | 1519 | 1643 | 1510 | 1400 | 1468 | 1500 | 1579 | 1575.1695 | 1629 | 1640 | 1588 | 1540 | 1540 | 1500 | 1320 | 1460 | 1465 | 1500.33 | 0
| 1 | 3 | 1580 | 1510 | 1539.5087 | 1568 | 1490 | 1475 | 1610 | 1570 | 1570 | 1420 | 1698 | 1570 | 1569 | 1600 | 1460 | 1500 | 1449 | 1400 | 1574 | 1581.8335 | 1561 | 1610 | 1562 | 1530 | 1520 | 1530 | 1330 | 1510 | 1485 | 1522.06 | 16
| 1 | 4 | 1540 | 1500 | 1540.5084 | 1565 | 1510 | 1495 | 1580 | 1560 | 1570 | 1400 | 1670 | 1600 | 1515 | 1595 | 1490 | 1400 | 1473 | 1600 | 1610 | 1557.3896 | 1588 | 1580 | 1545 | 1550 | 1490 | 1510 | 1350 | 1480 | 1481 | 1536.7 | 1533.4576
| 1 | 5 | 1500 | 1570 | 1529.2716 | 1638 | 1630 | 1515 | 1620 | 1520 | 1610 | 1470 | 1649 | 1550 | 1449 | 1629 | 1480 | 1400 | 1496 | 1600 | 1576 | 1555.2459 | 1565 | 1600 | 1495 | 1520 | 1530 | 1480 | 1320 | 1510 | 1454 | 1524.78 | 21
| 1 | 6 | 1500 | 1550 | 1542.4017 | 1606 | 1490 | 1540 | 1600 | 1560 | 1590 | 1450 | 1648 | 1530 | 1589 | 1614 | 1500 | 1400 | 1485 | 1600 | 1617 | 1561.8489 | 1525 | 1580 | 1534 | 1550 | 1510 | 1480 | 1310 | 1570 | 1463 | 1524.49
| Mean |  | 1493.3333 | 1525 | 1537.6832 | 1584.3333 | 1508.3333 | 1506.6667 | 1593.3333 | 1556.6667 | 1580 | 1436.6667 | 1649.6667 | 1536.6667 | 1526.6667 | 1615.1667 | 1485 | 1416.6667 | 1474.1667 | 1533.3333 | 1590.5 | 1567.5152 | 1578.6667 | 1600 | 1547.3333 | 1540 | 1521.6667 | 1498.3333 | 1325 | 1505 | 1468 | 1518.2383
| Median |  | 1500 | 1525 | 1540.0086 | 1574.5 | 1490 | 1505 | 1590 | 1560 | 1575 | 1440 | 1648.5 | 1540 | 1519 | 1612 | 1485 | 1400 | 1473.5 | 1550 | 1583 | 1567.7265 | 1576.5 | 1595 | 1552.5 | 1545 | 1525 | 1495 | 1320 | 1505 | 1464 | 1523.275
| Std Dev. |  | 65.3197 | 33.9116 | 8.7767 | 32.6231 | 64.0052 | 22.9492 | 19.6638 | 20.6559 | 17.8885 | 25.0333 | 31.7406 | 46.3321 | 48.9149 | 18.0601 | 18.7083 | 40.8248 | 15.9175 | 81.6497 | 18.4905 | 10.8245 | 36.4454 | 22.8035 | 31.4685 | 12.6491 | 19.4079 | 19.4079 | 13.784 | 37.2827 | 12.2638 | 14.5105
| Rel.Std.Dev. |  | 4.374% | 2.224% | 0.5708% | 2.059% | 4.243% | 1.523% | 1.234% | 1.327% | 1.132% | 1.742% | 1.924% | 3.015% | 3.204% | 1.118% | 1.26% | 2.882% | 1.08% | 5.325% | 1.163% | 0.6905% | 2.309% | 1.425% | 2.034% | 0.8214% | 1.275% | 1.295% | 1.04% | 2.477% | 0.8354% | 0.9557%
| PDM3 |  | -2.617% | -0.5515% | 0.2756% | 3.318% | -1.638% | -1.747% | 3.905% | 1.514% | 3.035% | -6.312% | 7.578% | 0.2093% | -0.4429% | 5.328% | -3.16% | -7.616% | -3.866% | -0.008103% | 3.72% | 2.221% | 2.948% | 4.339% | 0.9049% | 0.4266% | -0.7689% | -2.291% | -13.59% | -1.856% | -4.269% | -0.9925%
| Z-Score (Absolute) |  | 0.6 | 0.08 | 0.13 | 0.89 | 0.75 | 0.38 | 1.04 | 0.44 | 0.82 | 1.53 | 1.96 | 0.11 | 0.05 | 1.39 | 0.74 | 1.85 | 0.91 | 0.05 | 0.99 | 0.61 | 0.8 | 1.14 | 0.28 | 0.16 | 0.14 | 0.52 | 3.35 | 0.41 | 1.01 | 0.19
| Analytical results for Dy in OREAS 905 (Certified Value 3.72 ppm) | [31 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 05; 02: 09; 03: 10; 04: 12; 05: 13; 06: 17; 07: 26; 08: 30; 09: 33 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS | 2
3
| 1 | 1 | 01: 3.8; 02: 3.6; 03: 3.56; 04: 2.9; 05: 3.3; 06: 3.7; 07: 4.1; 08: 3.2; 09: 3.3 | 1
| 1 | 2 | 01: 3.9; 02: 3.7; 03: 3.6; 04: 2.8; 05: 3.2; 06: 3.83; 07: 3.7; 08: 3.2; 09: 4 | 0
| 1 | 3 | 01: 4; 02: 3.55; 03: 3.66; 04: 2.9; 05: 3.8; 06: 3.76; 07: 3.85; 08: 3.3; 09: 3.6 | 16
| 1 | 4 | 01: 4; 02: 3.8; 03: 3.5; 04: 2.9; 05: 3.6; 06: 3.73; 07: 3.8; 08: 3.4; 09: 3.4 | 3.7245
| 1 | 5 | 01: 4.1; 02: 3.95; 03: 3.59; 04: 2.3; 05: 3.4; 06: 3.73; 07: 3.8; 08: 3.1; 09: 3.8 | 22
| 1 | 6 | 01: 4.2; 02: 3.9; 03: 3.61; 04: 2.5; 05: 3.9; 06: 3.82; 07: 3.8; 08: 3.1; 09: 3.8
| Mean |  | 01: 4; 02: 3.75; 03: 3.5867; 04: 2.7167; 05: 3.5333; 06: 3.7617; 07: 3.8417; 08: 3.2167; 09: 3.65
| Median |  | 01: 4; 02: 3.75; 03: 3.595; 04: 2.85; 05: 3.5; 06: 3.745; 07: 3.8; 08: 3.2; 09: 3.7
| Std Dev. |  | 01: 0.14142; 02: 0.16125; 03: 0.053541; 04: 0.25626; 05: 0.28048; 06: 0.052694; 07: 0.13571; 08: 0.1169; 09: 0.26646
| Rel.Std.Dev. |  | 01: 3.536%; 02: 4.3%; 03: 1.493%; 04: 9.433%; 05: 7.938%; 06: 1.401%; 07: 3.533%; 08: 3.634%; 09: 7.3%
| PDM3 |  | 01: 7.396%; 02: 0.684%; 03: -3.701%; 04: -27.06%; 05: -5.133%; 06: 0.9973%; 07: 3.145%; 08: -13.64%; 09: -2.001%
| Z-Score (Absolute) |  | 01: 2.02; 02: 0.58; 03: 0.37; 04: 5.39; 05: 0.67; 06: 0.65; 07: 0.81; 08: 2.5; 09: 0
| Analytical results for Er in OREAS 905 (Certified Value 1.12 ppm) | [31 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 05; 02: 09; 03: 10; 04: 12; 05: 13; 06: 17; 07: 26; 08: 30; 09: 33 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS | 2
3
| 1 | 1 | 01: 1.1; 02: 1.1; 03: 1.15; 04: 0.8; 05: 1; 06: 1.14; 07: 1.15; 08: 1.9; 09: 1.1 | 1
| 1 | 2 | 01: 1.1; 02: 1.2; 03: 1.19; 04: 0.8; 05: 1; 06: 1.16; 07: 1.1; 08: 2; 09: 1.1 | 13
| 1 | 3 | 01: 1.1; 02: 1.1; 03: 1.16; 04: 0.8; 05: 1.1; 06: 1.13; 07: 1.1; 08: 1.9; 09: 1 | 16
| 1 | 4 | 01: 1.1; 02: 1.1; 03: 1.13; 04: 0.9; 05: 1.1; 06: 1.12; 07: 1.1; 08: 2; 09: 1.2 | 1.1164
| 1 | 5 | 01: 1.1; 02: 1.1; 03: 1.16; 04: 0.7; 05: 1.1; 06: 1.14; 07: 1.15; 08: 1.9; 09: 1.2 | 23
| 1 | 6 | 01: 1.1; 02: 1.15; 03: 1.15; 04: 0.7; 05: 1.1; 06: 1.11; 07: 1.1; 08: 1.9; 09: 1.1
| Mean |  | 01: 1.1; 02: 1.125; 03: 1.1567; 04: 0.78333; 05: 1.0667; 06: 1.1333; 07: 1.1167; 08: 1.9333; 09: 1.1167
| Median |  | 01: 1.1; 02: 1.1; 03: 1.155; 04: 0.8; 05: 1.1; 06: 1.135; 07: 1.1; 08: 1.9; 09: 1.1
| Std Dev. |  | 01: 2.4324e-16; 02: 0.041833; 03: 0.019664; 04: 0.075277; 05: 0.05164; 06: 0.017512; 07: 0.02582; 08: 0.05164; 09: 0.075277
| Rel.Std.Dev. |  | 01: 2.211e-14%; 02: 3.718%; 03: 1.7%; 04: 9.61%; 05: 4.841%; 06: 1.545%; 07: 2.312%; 08: 2.671%; 09: 6.741%
| PDM3 |  | 01: -1.472%; 02: 0.7678%; 03: 3.604%; 04: -29.84%; 05: -4.457%; 06: 1.514%; 07: 0.02133%; 08: 73.17%; 09: 0.02133%
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.34; 03: 1.62; 04: 13.49; 05: 2.02; 06: 0.67; 07: 0; 08: 33.04; 09: 0
| Analytical results for Eu in OREAS 905 (Certified Value 1.42 ppm) | [31 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 05; 02: 09; 03: 10; 04: 12; 05: 13; 06: 17; 07: 20; 08: 26; 09: 30; 10: 33 | ppm
|  |  | 01: 4A*MS; 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*OES/MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*OES/MS | 2
3
| 1 | 1 | 01: 1.48; 02: 1.45; 03: 2.34; 04: 1.34; 05: 1.3; 06: 1.47; 07: 1.5046; 08: 1.4; 09: 2.74; 10: 1.3 | 1
| 1 | 2 | 01: 1.56; 02: 1.6; 03: 2.34; 04: 1.4; 05: 1.3; 06: 1.4; 07: 1.545; 08: 1.35; 09: 2.78; 10: 1.4 | 0
| 1 | 3 | 01: 1.55; 02: 1.45; 03: 2.27; 04: 1.26; 05: 1.5; 06: 1.36; 07: 1.5336; 08: 1.4; 09: 2.87; 10: 1.4 | 16
| 1 | 4 | 01: 1.6; 02: 1.5; 03: 2.2; 04: 1.35; 05: 1.4; 06: 1.42; 07: 1.5688; 08: 1.4; 09: 2.94; 10: 1.2 | 1.4225
| 1 | 5 | 01: 1.58; 02: 1.55; 03: 2.29; 04: 1.02; 05: 1.3; 06: 1.41; 07: 1.604; 08: 1.35; 09: 2.63; 10: 1.3 | 24
| 1 | 6 | 01: 1.55; 02: 1.55; 03: 2.29; 04: 1.27; 05: 1.5; 06: 1.37; 07: 1.4819; 08: 1.35; 09: 2.62; 10: 1.1
| Mean |  | 01: 1.5533; 02: 1.5167; 03: 2.2883; 04: 1.2733; 05: 1.3833; 06: 1.405; 07: 1.5397; 08: 1.375; 09: 2.7633; 10: 1.2833
| Median |  | 01: 1.555; 02: 1.525; 03: 2.29; 04: 1.305; 05: 1.35; 06: 1.405; 07: 1.5393; 08: 1.375; 09: 2.76; 10: 1.3
| Std Dev. |  | 01: 0.040825; 02: 0.060553; 03: 0.051929; 04: 0.13471; 05: 0.098319; 06: 0.03937; 07: 0.043857; 08: 0.027386; 09: 0.12785; 10: 0.1169
| Rel.Std.Dev. |  | 01: 2.628%; 02: 3.993%; 03: 2.269%; 04: 10.58%; 05: 7.107%; 06: 2.802%; 07: 2.848%; 08: 1.992%; 09: 4.627%; 10: 9.109%
| PDM3 |  | 01: 9.194%; 02: 6.617%; 03: 60.86%; 04: -10.49%; 05: -2.756%; 06: -1.233%; 07: 8.233%; 08: -3.342%; 09: 94.25%; 10: -9.786%
| Z-Score (Absolute) |  | 01: 0.7; 02: 0.42; 03: 6.26; 04: 1.03; 05: 0.59; 06: 0.42; 07: 0.6; 08: 0.65; 09: 9.85; 10: 1.34
| Analytical results for Fe in OREAS 905 (Certified Value 4.08 wt.%) | [31 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 3.79 | 4.3678 | 4.01 | 4.21 | 4.03 | 4.04 | 4.28 | 4.26 | 3.86 | 4.09 | 4.1 | 3.85 | 4.32 | 3.96 | 4.05 | 4.06 | 4.21 | 4.086 | 4.0962 | 4.31 | 3.7 | 4.09 | 4.08 | 3.92 | 4.08 | 3.94 | 3.89 | 4.17 |  |  | 1
| 1 | 2 | 3.87 | 4.2586 | 4 | 4.34 | 4 | 4.05 | 4.4 | 4.27 | 3.85 | 4.4 | 3.95 | 3.78 | 4.43 | 3.84 | 4.06 | 4.06 | 4.29 | 4.067 | 4.1247 | 4.37 | 3.72 | 4.21 | 4.09 | 3.94 | 4.12 | 3.95 | 3.85 | 4.17 |  |  | 0
| 1 | 3 | 4.2 | 4.2633 | 4.14 | 4.22 | 3.91 | 4.08 | 4.36 | 4.31 | 3.8 | 4.08 | 4.23 | 4.17 | 4.28 | 4.17 | 4.08 | 3.99 | 3.91 | 3.99 | 4.0486 | 4.14 | 3.71 | 4.09 | 4.08 | 4.03 | 4.15 | 4.04 | 3.71 | 4.21 |  |  | 16
| 1 | 4 | 4.07 | 4.2162 | 4.02 | 4.27 | 4 | 4.04 | 4.31 | 4.2 | 3.68 | 4.03 | 4.27 | 4.05 | 4.36 | 3.9 | 3.84 | 4.05 | 4.11 | 4.143 | 4.0703 | 4.26 | 3.74 | 4.1 | 4.03 | 3.95 | 4.21 | 3.98 | 3.92 | 4.27 |  |  | 4.0764
| 1 | 5 | 4.04 | 4.3559 | 4.07 | 4.21 | 4.02 | 4.13 | 4.43 | 4.43 | 3.84 | 4.01 | 4.1 | 3.86 | 4.44 | 3.91 | 3.91 | 4.09 | 4.13 | 4.133 | 4.0154 | 4.14 | 3.74 | 4.02 | 4.11 | 3.89 | 4.12 | 3.97 | 3.74 | 4.21 |  |  | 25
| 1 | 6 | 4.04 | 4.3709 | 3.94 | 4.4 | 4.09 | 4.11 | 4.29 | 4.39 | 3.92 | 4.02 | 4.15 | 4.18 | 4.47 | 3.83 | 3.81 | 4.05 | 4.17 | 4.057 | 3.9684 | 4.07 | 3.67 | 4.08 | 4.06 | 3.89 | 4.07 | 4.07 | 3.73 | 4.21
| Mean |  | 4.0017 | 4.3055 | 4.03 | 4.275 | 4.0083 | 4.075 | 4.345 | 4.31 | 3.825 | 4.105 | 4.1333 | 3.9817 | 4.3833 | 3.935 | 3.9583 | 4.05 | 4.1367 | 4.0793 | 4.0539 | 4.215 | 3.7133 | 4.0983 | 4.075 | 3.9367 | 4.125 | 3.9917 | 3.8067 | 4.2067
| Median |  | 4.04 | 4.3096 | 4.015 | 4.245 | 4.01 | 4.065 | 4.335 | 4.29 | 3.845 | 4.055 | 4.125 | 3.955 | 4.395 | 3.905 | 3.98 | 4.055 | 4.15 | 4.0765 | 4.0595 | 4.2 | 3.715 | 4.09 | 4.08 | 3.93 | 4.12 | 3.975 | 3.795 | 4.21
| Std Dev. |  | 0.1477 | 0.067312 | 0.068118 | 0.079183 | 0.058452 | 0.038341 | 0.061563 | 0.086023 | 0.080932 | 0.14816 | 0.11325 | 0.17452 | 0.075011 | 0.1247 | 0.1199 | 0.032863 | 0.12817 | 0.055874 | 0.056356 | 0.11606 | 0.026583 | 0.061779 | 0.027386 | 0.052026 | 0.050892 | 0.051929 | 0.090921 | 0.036697
| Rel.Std.Dev. |  | 3.691% | 1.563% | 1.69% | 1.852% | 1.458% | 0.9409% | 1.417% | 1.996% | 2.116% | 3.609% | 2.74% | 4.383% | 1.711% | 3.169% | 3.029% | 0.8114% | 3.098% | 1.37% | 1.39% | 2.754% | 0.7159% | 1.507% | 0.6721% | 1.322% | 1.234% | 1.301% | 2.388% | 0.8724%
| PDM3 |  | -1.834% | 5.619% | -1.139% | 4.872% | -1.67% | -0.03464% | 6.589% | 5.73% | -6.167% | 0.7013% | 1.396% | -2.324% | 7.529% | -3.469% | -2.897% | -0.6479% | 1.478% | 0.07166% | -0.5515% | 3.4% | -8.907% | 0.5378% | -0.03464% | -3.428% | 1.192% | -2.079% | -6.617% | 3.195%
| Z-Score (Absolute) |  | 0.54 | 2.09 | 0.3 | 1.82 | 0.49 | 0.09 | 2.43 | 2.13 | 1.82 | 0.16 | 0.6 | 0.72 | 2.76 | 1.53 | 0.92 | 0.13 | 0.63 | 0.13 | 0.09 | 1.3 | 3.04 | 0.29 | 0.09 | 1.11 | 0.52 | 0.63 | 2.23 | 1.23
| Analytical results for Ga in OREAS 905 (Certified Value 25.1 ppm) | [31 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 25 | 26.4023 | 24.1 | 26.2 | 24.3 | 28.7 | 20 | 30.4 | 25.4 | 25 | 26.1 | 24.8 | 24.5 | 34 | 24.79 | 25 | 23.5 | 26.691 | 31 | 24.4 | 24.9 | 25.3 | < 0.2 | 23.9 | 26.2 | 23.8 |  |  |  |  | 1
| 1 | 2 | 24.9 | 25.6953 | 25.4 | 26.6 | 24.2 | 29 | 20.9 | 31.8 | 24.9 | 26 | 26.9 | 25 | 25.5 | 35 | 24.38 | 26 | 22.7 | 26.9923 | 32 | 24.4 | 22.7 | 26.2 | < 0.2 | 23.9 | 26.1 | 24.32 |  |  |  |  | 0
| 1 | 3 | 25.9 | 26.0233 | 24.3 | 26.5 | 23.3 | 29.1 | 20.6 | 30.8 | 24.6 | 25 | 25.6 | 25.9 | 25.3 | 32 | 25.13 | 24 | 22.7 | 26.7082 | 31 | 23.8 | 25.3 | 26.1 | < 0.2 | 24.5 | 26.6 | 25.05 |  |  |  |  | 16
| 1 | 4 | 24.4 | 25.9192 | 24.9 | 27.2 | 23.9 | 28 | 20.5 | 31.2 | 23.8 | 25 | 25.3 | 25.7 | 25.6 | 29 | 23.67 | 25 | 23.5 | 26.0403 | 32 | 24.4 | 26 | 25.9 | < 0.2 | 24.1 | 26.3 | 24.64 |  |  |  |  | 25.118
| 1 | 5 | 24.5 | 25.883 | 23.8 | 27 | 24 | 28.2 | 20.4 | 32 | 24.8 | 25 | 25.1 | 23.9 | 26.1 | 33 | 25.14 | 26 | 22.7 | 26.2939 | 31 | 24 | 24.7 | 26.5 | < 0.2 | 24.3 | 25.9 | 23.76 |  |  |  |  | 26
| 1 | 6 | 24.5 | 26.1942 | 24.2 | 27 | 23.5 | 30 | 20.5 | 31.2 | 24.8 | 26 | 25.6 | 26.1 | 25.5 | 33 | 24.41 | 25 | 23.5 | 26.1959 | 31 | 25 | 25.3 | 24.9 | < 0.2 | 24.9 | 25.5 | 25.29
| Mean |  | 24.8667 | 26.0195 | 24.45 | 26.75 | 23.8667 | 28.8333 | 20.4833 | 31.2333 | 24.7167 | 25.3333 | 25.7667 | 25.2333 | 25.4167 | 32.6667 | 24.5867 | 25.1667 | 23.1 | 26.4869 | 31.3333 | 24.3333 | 24.8167 | 25.8167 |  | 24.2667 | 26.1 | 24.4767
| Median |  | 24.7 | 25.9712 | 24.25 | 26.8 | 23.95 | 28.85 | 20.5 | 31.2 | 24.8 | 25 | 25.6 | 25.35 | 25.5 | 33 | 24.6 | 25 | 23.1 | 26.4925 | 31 | 24.4 | 25.1 | 26 |  | 24.2 | 26.15 | 24.48
| Std Dev. |  | 0.56095 | 0.24933 | 0.58907 | 0.37815 | 0.39328 | 0.7174 | 0.29269 | 0.59889 | 0.52313 | 0.5164 | 0.65013 | 0.82865 | 0.52313 | 2.0656 | 0.55802 | 0.75277 | 0.43818 | 0.36536 | 0.5164 | 0.41312 | 1.1286 | 0.60139 |  | 0.38816 | 0.37417 | 0.6345
| Rel.Std.Dev. |  | 2.256% | 0.9582% | 2.409% | 1.414% | 1.648% | 2.488% | 1.429% | 1.917% | 2.117% | 2.038% | 2.523% | 3.284% | 2.058% | 6.323% | 2.27% | 2.991% | 1.897% | 1.379% | 1.648% | 1.698% | 4.548% | 2.329% |  | 1.6% | 1.434% | 2.592%
| PDM3 |  | -1.001% | 3.589% | -2.659% | 6.497% | -4.982% | 14.79% | -18.45% | 24.35% | -1.598% | 0.8573% | 2.583% | 0.4592% | 1.189% | 30.05% | -2.115% | 0.1938% | -8.034% | 5.45% | 24.74% | -3.124% | -1.2% | 2.782% |  | -3.389% | 3.91% | -2.553%
| Z-Score (Absolute) |  | 0.32 | 0.67 | 0.67 | 1.29 | 1.17 | 3.07 | 4.06 | 5.12 | 0.29 | 0.08 | 0.45 | 0.01 | 0.31 | 6.34 | 0.56 | 0.06 | 1.83 | 1.06 | 5.2 | 0.77 | 0 | 0.49 | NA | 0.83 | 0.73 | 0.65
| Analytical results for Gd in OREAS 905 (Certified Value 5.9 ppm) | [31 empty] | Certified
Gd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 17; 08: 26; 09: 30; 10: 33 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 01: 7; 03: 5.6; 04: 5.9; 05: 5.4; 06: 5.5; 07: 6.23; 08: 6.4; 09: 11; 10: 6.2 |  |  |  |  | 1
| 1 | 2 | 01: 7; 03: 6.2; 04: 5.88; 05: 5.1; 06: 5.4; 07: 6.46; 08: 6; 09: 11.4; 10: 5.9 |  |  |  |  | 0
| 1 | 3 | 01: 7.1; 03: 5.6; 04: 5.89; 05: 5.3; 06: 6.1; 07: 6.08; 08: 6.2; 09: 11.5; 10: 5.6 |  |  |  |  | 16
| 1 | 4 | 01: 7.2; 03: 5.6; 04: 5.7; 05: 5.8; 06: 6; 07: 6.37; 08: 6.4; 09: 12.1; 10: 5.9 |  |  |  |  | 5.8982
| 1 | 5 | 01: 7.3; 03: 6; 04: 5.92; 05: 4.3; 06: 5.8; 07: 6.44; 08: 6.2; 09: 11; 10: 5.7 |  |  |  |  | 27
| 1 | 6 | 01: 7.7; 03: 5.8; 04: 5.99; 05: 5; 06: 6.3; 07: 6.33; 08: 6; 09: 11.3; 10: 6
| Mean |  | 01: 7.2167; 03: 5.8; 04: 5.88; 05: 5.15; 06: 5.85; 07: 6.3183; 08: 6.2; 09: 11.3833; 10: 5.8833
| Median |  | 01: 7.15; 03: 5.7; 04: 5.895; 05: 5.2; 06: 5.9; 07: 6.35; 08: 6.2; 09: 11.35; 10: 5.9
| Std Dev. |  | 01: 0.26394; 03: 0.25298; 04: 0.09654; 05: 0.501; 06: 0.35071; 07: 0.14303; 08: 0.17889; 09: 0.40702; 10: 0.2137
| Rel.Std.Dev. |  | 01: 3.657%; 03: 4.362%; 04: 1.642%; 05: 9.728%; 06: 5.995%; 07: 2.264%; 08: 2.885%; 09: 3.576%; 10: 3.632%
| PDM3 |  | 01: 22.35%; 03: -1.666%; 04: -0.3092%; 05: -12.69%; 06: -0.8178%; 07: 7.122%; 08: 5.116%; 09: 93%; 10: -0.2527%
| Z-Score (Absolute) |  | 01: 3.09; 03: 0.28; 04: 0; 05: 1.42; 06: 0.16; 07: 0.96; 08: 0.67; 09: 12.98; 10: 0.08
| Analytical results for Ge in OREAS 905 (Indicative Value 0.54 ppm) | [31 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 15: Lab; 16: Lab; 17: Lab |  |  |  |  | 1
| No | No | 01: 01; 03: 03; 04: 04; 05: 06; 06: 07; 07: 08; 08: 10; 09: 11; 10: 12; 11: 13; 12: 17; 13: 24; 15: 28; 16: 29; 17: 31 |  |  |  |  | ppm
|  |  | 01: 4A*OES/MS; 03: 4A*OES; 04: 4A*OES/MS; 05: 4A*OES/MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*OES; 10: 4A*MS; 11: 4A*MS; 12: 4A*MS; 13: 4A*OES/MS; 15: 4A*OES/MS; 16: 4A*OES/MS; 17: 4A*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 01: 0.25; 03: 1.3298; 04: 0.4; 05: 0.17; 06: 0.18; 07: 0.5; 08: < 0.1; 09: < 10; 10: 0.6; 11: 0.4; 12: 1.25; 13: 4.5; 15: 0.19; 16: 0.15; 17: 0.22 |  |  |  |  | 1
| 1 | 2 | 01: 0.31; 03: 1.2135; 04: 0.6; 05: 0.17; 06: 0.18; 07: 0.5; 08: < 0.1; 09: < 10; 10: 0.6; 11: < 0.1; 12: 1.45; 13: 3.9; 15: 0.17; 16: 0.2; 17: 0.22 |  |  |  |  | 4
| 1 | 3 | 01: 0.27; 03: 1.2932; 04: 0.4; 05: 0.17; 06: 0.19; 07: 0.5; 08: < 0.1; 09: < 10; 10: 0.6; 11: 1.1; 12: 1.33; 13: 4; 15: 0.3; 16: 0.19; 17: 0.19 |  |  |  |  | 16
| 1 | 4 | 01: 0.27; 03: 1.3139; 04: 0.5; 05: 0.17; 06: 0.18; 07: 0.5; 08: < 0.1; 09: < 10; 10: 0.6; 11: 1; 12: 1.39; 13: 4.2; 15: 0.34; 16: 0.22; 17: 0.17 |  |  |  |  | 0.53686
| 1 | 5 | 01: 0.28; 03: 1.4184; 04: 0.5; 05: 0.17; 06: 0.18; 07: 0.5; 08: < 0.1; 09: < 10; 10: 0.4; 11: 0.9; 12: 1.4; 13: 4; 15: 0.23; 16: 0.19; 17: 0.21 |  |  |  |  | 10
| 1 | 6 | 01: 0.27; 03: 1.3169; 04: 0.4; 05: 0.16; 06: 0.19; 07: 0.5; 08: < 0.1; 09: < 10; 10: 0.5; 11: 0.9; 12: 1.44; 13: 4.2; 15: 0.22; 16: 0.21; 17: 0.22
| Mean |  | 01: 0.275; 03: 1.3143; 04: 0.46667; 05: 0.16833; 06: 0.18333; 07: 0.5; 10: 0.55; 11: 0.86; 12: 1.3767; 13: 4.1333; 15: 0.24167; 16: 0.19333; 17: 0.205
| Median |  | 01: 0.27; 03: 1.3154; 04: 0.45; 05: 0.17; 06: 0.18; 07: 0.5; 10: 0.6; 11: 0.9; 12: 1.395; 13: 4.1; 15: 0.225; 16: 0.195; 17: 0.215
| Std Dev. |  | 01: 0.019748; 03: 0.065881; 04: 0.08165; 05: 0.0040825; 06: 0.005164; 07: 0; 10: 0.083666; 11: 0.27019; 12: 0.075277; 13: 0.21602; 15: 0.065549; 16: 0.024221; 17: 0.020736
| Rel.Std.Dev. |  | 01: 7.181%; 03: 5.013%; 04: 17.5%; 05: 2.425%; 06: 2.817%; 07: 0%; 10: 15.21%; 11: 31.42%; 12: 5.468%; 13: 5.226%; 15: 27.12%; 16: 12.53%; 17: 10.12%
| PDM3 |  | 01: -48.78%; 03: 144.8%; 04: -13.07%; 05: -68.64%; 06: -65.85%; 07: -6.865%; 10: 2.448%; 11: 60.19%; 12: 156.4%; 13: 669.9%; 15: -54.98%; 16: -63.99%; 17: -61.81%
| Z-Score (Absolute) |  | 01: 0.49; 03: 2.09; 04: 0; 05: 0.74; 06: 0.7; 07: 0.08; 08: NA; 09: NA; 10: 0.21; 11: 1.25; 12: 2.24; 13: 9.05; 15: 0.56; 16: 0.67; 17: 0.65
| Analytical results for Hf in OREAS 905 (Certified Value 6.84 ppm) | [31 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 17 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [9 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [9 empty] | 2
3
| 1 | 1 | 6.3 | 6.6279 | 7.12 | 6.47 | 6.7 | 6.7 | 6.7 | 6.2 | 6.96 | 3.9 | 7.1 | 6.84 | 7.51 | 7.9 | 7.2 | 6.8 | 6.2 | 7.12 | 7.1 | 6.97 | 6.99 | [9 empty] | 1
| 1 | 2 | 6.1 | 6.7628 | 7.11 | 6.67 | 6.5 | 6.6 | 6.8 | 6.8 | 7.02 | 6.3 | 7.4 | 6.83 | 7.39 | 7.4 | 7 | 6 | 6.4 | 6.82 | 6.8 | 7.05 | 7.02 | [9 empty] | 0
| 1 | 3 | 6.5 | 6.6645 | 7.1 | 6.53 | 6.5 | 6.8 | 6.8 | 6.2 | 7.08 | 6.2 | 7.8 | 6.52 | 7.08 | 7.3 | 7 | 6.9 | 6.4 | 7.12 | 7.1 | 6.97 | 7.02 | [9 empty] | 16
| 1 | 4 | 6.3 | 6.7825 | 7.05 | 6.67 | 6.6 | 6.7 | 6.5 | 6.4 | 7.04 | 6.1 | 7.6 | 6.46 | 7.29 | 7.1 | 6.6 | 6.9 | 6.5 | 7.07 | 6.9 | 7.09 | 7.08 | [9 empty] | 6.8444
| 1 | 5 | 6.4 | 6.5832 | 7.21 | 6.63 | 6.5 | 6.8 | 6.7 | 6.4 | 7.22 | 6.4 | 7.3 | 6.57 | 7.3 | 7.4 | 7.2 | 6.6 | 6.4 | 6.91 | 7.1 | 7 | 7.18 | [9 empty] | 28
| 1 | 6 | 6.4 | 6.8448 | 7 | 6.8 | 6.5 | 6.9 | 6.7 | 6.4 | 7.13 | 6.5 | 7.9 | 6.7 | 7.59 | 7.5 | 7 | 6.8 | 6.1 | 6.48 | 7.1 | 6.8 | 7.21
| Mean |  | 6.3333 | 6.711 | 7.0983 | 6.6283 | 6.55 | 6.75 | 6.7 | 6.4 | 7.075 | 5.9 | 7.5167 | 6.6533 | 7.36 | 7.4333 | 7 | 6.6667 | 6.3333 | 6.92 | 7.0167 | 6.98 | 7.0833
| Median |  | 6.35 | 6.7136 | 7.105 | 6.65 | 6.5 | 6.75 | 6.7 | 6.4 | 7.06 | 6.25 | 7.5 | 6.635 | 7.345 | 7.4 | 7 | 6.8 | 6.4 | 6.99 | 7.1 | 6.985 | 7.05
| Std Dev. |  | 0.13663 | 0.10109 | 0.070828 | 0.11635 | 0.083666 | 0.10488 | 0.10954 | 0.21909 | 0.091159 | 0.98995 | 0.30605 | 0.16145 | 0.18067 | 0.26583 | 0.21909 | 0.34448 | 0.15055 | 0.24747 | 0.13292 | 0.1 | 0.091797
| Rel.Std.Dev. |  | 2.157% | 1.506% | 0.9978% | 1.755% | 1.277% | 1.554% | 1.635% | 3.423% | 1.288% | 16.78% | 4.072% | 2.427% | 2.455% | 3.576% | 3.13% | 5.167% | 2.377% | 3.576% | 1.894% | 1.433% | 1.296%
| PDM3 |  | -7.467% | -1.95% | 3.71% | -3.157% | -4.301% | -1.379% | -2.11% | -6.493% | 3.369% | -13.8% | 9.822% | -2.792% | 7.533% | 8.604% | 2.273% | -2.597% | -7.467% | 1.104% | 2.517% | 1.981% | 3.491%
| Z-Score (Absolute) |  | 1.21 | 0.23 | 0.77 | 0.45 | 0.65 | 0.13 | 0.26 | 1.04 | 0.71 | 1.3 | 1.86 | 0.38 | 1.45 | 1.64 | 0.52 | 0 | 1.09 | 0.31 | 0.67 | 0.47 | 0.73
| Analytical results for Hg in OREAS 905 (Indicative Value 0.027 ppm) | [31 empty] | Indicative
Hg | Round | Replicate | 01: Lab |  |  |  |  | 1
| No | No | 01: 12 |  |  |  |  | ppm
|  |  | 01: 4A*MS |  |  |  |  | 3
3
| 1 | 1 | 01: 0.03 |  |  |  |  | 1
| 1 | 2 | 01: 0.03 |  |  |  |  | 3
| 1 | 3 | 01: 0.04 |  |  |  |  | 16
| 1 | 4 | 01: 0.03 |  |  |  |  | 0.026667
| 1 | 5 | 01: 0.02 |  |  |  |  | 9
| 1 | 6 | 01: 0.01
| Mean |  | 01: 0.026667
| Median |  | 01: 0.03
| Std Dev. |  | 01: 0.010328
| Rel.Std.Dev. |  | 01: 38.73%
| PDM3 |  | 01: -1.221e-13%
Z-Score (Absolute) | bar: 01
| Analytical results for Ho in OREAS 905 (Certified Value 0.5 ppm) | [31 empty] | Certified
Ho | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 17; 08: 26; 09: 30; 10: 33 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 01: 0.47; 03: 0.5; 04: 0.53; 05: 0.5; 06: 0.5; 07: 0.54; 08: 0.54; 09: 0.42; 10: 0.4 |  |  |  |  | 1
| 1 | 2 | 01: 0.49; 03: 0.6; 04: 0.53; 05: 0.5; 06: 0.4; 07: 0.54; 08: 0.5; 09: 0.43; 10: 0.6 |  |  |  |  | 0
| 1 | 3 | 01: 0.48; 03: 0.52; 04: 0.53; 05: 0.5; 06: 0.5; 07: 0.54; 08: 0.52; 09: 0.44; 10: 0.5 |  |  |  |  | 16
| 1 | 4 | 01: 0.51; 03: 0.52; 04: 0.51; 05: 0.5; 06: 0.5; 07: 0.53; 08: 0.52; 09: 0.44; 10: 0.6 |  |  |  |  | 0.4987
| 1 | 5 | 01: 0.49; 03: 0.54; 04: 0.53; 05: 0.4; 06: 0.5; 07: 0.54; 08: 0.52; 09: 0.43; 10: 0.5 |  |  |  |  | 29
| 1 | 6 | 01: 0.49; 03: 0.52; 04: 0.54; 05: 0.4; 06: 0.5; 07: 0.52; 08: 0.52; 09: 0.42; 10: 0.5
| Mean |  | 01: 0.48833; 03: 0.53333; 04: 0.52833; 05: 0.46667; 06: 0.48333; 07: 0.535; 08: 0.52; 09: 0.43; 10: 0.51667
| Median |  | 01: 0.49; 03: 0.52; 04: 0.53; 05: 0.5; 06: 0.5; 07: 0.54; 08: 0.52; 09: 0.43; 10: 0.5
| Std Dev. |  | 01: 0.013292; 03: 0.035024; 04: 0.0098319; 05: 0.05164; 06: 0.040825; 07: 0.0083666; 08: 0.012649; 09: 0.0089443; 10: 0.075277
| Rel.Std.Dev. |  | 01: 2.722%; 03: 6.567%; 04: 1.861%; 05: 11.07%; 06: 8.447%; 07: 1.564%; 08: 2.433%; 09: 2.08%; 10: 14.57%
| PDM3 |  | 01: -2.079%; 03: 6.944%; 04: 5.941%; 05: -6.424%; 06: -3.082%; 07: 7.278%; 08: 4.27%; 09: -13.78%; 10: 3.602%
| Z-Score (Absolute) |  | 01: 1.04; 03: 0.12; 04: 0.43; 05: 1.84; 06: 1.23; 07: 0.67; 08: 0.12; 09: 3.19; 10: 0
| Analytical results for In in OREAS 905 (Certified Value 0.64 ppm) | [31 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 0.592 | 0.47165 | 0.63 | 0.66 | 0.636 | 0.659 | 0.55 | 0.6 | 0.82 | < 10 | 0.5 | 0.65 | 0.67 | 0.64 | 0.70582 | 1.02 | 0.68 | 0.641 | 0.58 | 0.37 | 0.64 | 0.65 | 0.63 | [7 empty] | 1
| 1 | 2 | 0.577 | 0.46624 | 0.64 | 0.66 | 0.638 | 0.658 | 0.55 | 0.7 | 0.83 | < 10 | 0.6 | 0.69 | 0.6 | 0.64 | 0.74983 | 0.8 | 0.62 | 0.574 | 0.595 | 0.4 | 0.614 | 0.65 | 0.71 | [7 empty] | 0
| 1 | 3 | 0.598 | 0.4904 | 0.63 | 0.64 | 0.615 | 0.66 | 0.54 | 0.6 | 0.79 | < 10 | 0.5 | 0.68 | 0.6 | 0.61 | 0.70169 | 0.7 | 0.6 | 0.615 | 0.605 | 0.4 | 0.642 | 0.67 | 0.67 | [7 empty] | 16
| 1 | 4 | 0.574 | 0.44212 | 0.63 | 0.69 | 0.644 | 0.665 | 0.54 | 0.6 | 0.79 | < 10 | 0.6 | 0.68 | 0.61 | 0.66 | 0.74281 | 0.73 | 0.64 | 0.651 | 0.596 | 0.4 | 0.63 | 0.65 | 0.68 | [7 empty] | 0.63586
| 1 | 5 | 0.579 | 0.44593 | 0.62 | 0.66 | 0.649 | 0.676 | 0.56 | 0.65 | 0.81 | < 10 | 0.6 | 0.64 | 0.62 | 0.65 | 0.70382 | 0.66 | 0.62 | 0.636 | 0.595 | 0.39 | 0.638 | 0.65 | 0.72 | [7 empty] | 30
| 1 | 6 | 0.583 | 0.47715 | 0.62 | 0.65 | 0.629 | 0.66 | 0.56 | 0.65 | 0.81 | < 10 | 0.6 | 0.69 | 0.6 | 0.62 | 0.75183 | 0.66 | 0.66 | 0.643 | 0.572 | 0.39 | 0.644 | 0.64 | 0.66
| Mean |  | 0.58383 | 0.46558 | 0.62833 | 0.66 | 0.63517 | 0.663 | 0.55 | 0.63333 | 0.80833 |  | 0.56667 | 0.67167 | 0.61667 | 0.63667 | 0.72597 | 0.76167 | 0.63667 | 0.62667 | 0.5905 | 0.39167 | 0.63467 | 0.65167 | 0.67833
| Median |  | 0.581 | 0.46894 | 0.63 | 0.66 | 0.637 | 0.66 | 0.55 | 0.625 | 0.81 |  | 0.6 | 0.68 | 0.605 | 0.64 | 0.72432 | 0.715 | 0.63 | 0.6385 | 0.595 | 0.395 | 0.639 | 0.65 | 0.675
| Std Dev. |  | 0.0093256 | 0.018568 | 0.0075277 | 0.016733 | 0.012024 | 0.0068118 | 0.0089443 | 0.040825 | 0.016021 |  | 0.05164 | 0.02137 | 0.027325 | 0.018619 | 0.024526 | 0.13688 | 0.029439 | 0.0285 | 0.012112 | 0.01169 | 0.011219 | 0.0098319 | 0.033116
| Rel.Std.Dev. |  | 1.597% | 3.988% | 1.198% | 2.535% | 1.893% | 1.027% | 1.626% | 6.446% | 1.982% |  | 9.113% | 3.182% | 4.431% | 2.924% | 3.378% | 17.97% | 4.624% | 4.548% | 2.051% | 2.985% | 1.768% | 1.509% | 4.882%
| PDM3 |  | -8.182% | -26.78% | -1.183% | 3.797% | -0.1087% | 4.269% | -13.5% | -0.397% | 27.12% |  | -10.88% | 5.632% | -3.018% | 0.1272% | 14.17% | 19.79% | 0.1272% | -1.445% | -7.133% | -38.4% | -0.1874% | 2.486% | 6.68%
| Z-Score (Absolute) |  | 1.09 | 3.51 | 0.17 | 0.48 | 0.03 | 0.54 | 1.78 | 0.07 | 3.53 | NA | 1.44 | 0.72 | 0.63 | 0 | 1.83 | 1.04 | 0 | 0.01 | 0.95 | 5.03 | 0.04 | 0.31 | 0.86
| Analytical results for K in OREAS 905 (Certified Value 2.88 wt.%) | [31 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.78 | 2.8481 | 2.82 | 2.98 | 2.86 | 2.85 | 2.92 | 2.97 | 2.93 | 2.84 | 2.92 | 2.3216 | 2.61 | 2.87 | 2.82 | 2.7235 | 2.46 | 2.725 | 2.8627 | 3.04 | 3.02 | 2.14 | 2.92 | 2.87 | 2.82 | 3.02 | 2.87 | 2.81 | 3.05 |  | 1
| 1 | 2 | 2.81 | 2.8313 | 2.81 | 3.03 | 2.84 | 2.86 | 2.99 | 2.93 | 2.93 | 2.89 | 2.72 | 2.4905 | 2.95 | 2.86 | 2.95 | 2.7434 | 2.4 | 2.794 | 2.9204 | 3.06 | 3.1 | 2.88 | 2.98 | 2.87 | 2.82 | 3.05 | 2.83 | 2.81 | 3.09 |  | 0
| 1 | 3 | 3.04 | 2.8309 | 2.76 | 2.95 | 2.79 | 2.87 | 2.97 | 2.95 | 2.78 | 2.74 | 3 | 2.8579 | 2.64 | 2.88 | 2.87 | 2.6939 | 2.33 | 2.814 | 2.8705 | 2.06 | 3.07 | 2.16 | 2.94 | 2.85 | 2.91 | 3.1 | 2.9 | 2.73 | 3.11 |  | 16
| 1 | 4 | 3 | 2.8494 | 2.72 | 3.04 | 2.85 | 2.83 | 2.93 | 2.9 | 2.62 | 2.69 | 2.77 | 2.7633 | 2.85 | 2.86 | 2.75 | 2.6998 | 2.4 | 2.873 | 2.8928 | 2.32 | 3 | 1.99 | 2.98 | 2.8 | 2.84 | 3.11 | 2.9 | 2.85 | 3.11 |  | 2.8752
| 1 | 5 | 2.96 | 2.8396 | 2.76 | 3.33 | 2.85 | 2.89 | 3.02 | 3.05 | 2.8 | 2.89 | 2.96 | 2.61 | 3 | 2.9 | 2.74 | 2.7313 | 2.47 | 2.686 | 2.8589 | 2.38 | 3.07 | 2.52 | 2.92 | 2.9 | 2.81 | 3.04 | 2.9 | 2.77 | 3.08 |  | 31
| 1 | 6 | 2.93 | 2.8617 | 2.71 | 3.11 | 2.91 | 2.86 | 2.93 | 3.03 | 2.95 | 2.87 | 2.61 | 2.9618 | 2.94 | 2.87 | 2.69 | 2.7419 | 2.54 | 2.931 | 2.8927 | 3 | 3.05 | 2.24 | 2.95 | 2.84 | 2.79 | 3.03 | 3 | 2.73 | 3.12
| Mean |  | 2.92 | 2.8435 | 2.7633 | 3.0733 | 2.85 | 2.86 | 2.96 | 2.9717 | 2.835 | 2.82 | 2.83 | 2.6675 | 2.8317 | 2.8733 | 2.8033 | 2.7223 | 2.4333 | 2.8038 | 2.883 | 2.6433 | 3.0517 | 2.3217 | 2.9483 | 2.855 | 2.8317 | 3.0583 | 2.9 | 2.7833 | 3.0933
| Median |  | 2.945 | 2.8439 | 2.76 | 3.035 | 2.85 | 2.86 | 2.95 | 2.96 | 2.865 | 2.855 | 2.845 | 2.6867 | 2.895 | 2.87 | 2.785 | 2.7274 | 2.43 | 2.804 | 2.8816 | 2.69 | 3.06 | 2.2 | 2.945 | 2.86 | 2.82 | 3.045 | 2.9 | 2.79 | 3.1
| Std Dev. |  | 0.10412 | 0.011914 | 0.045019 | 0.13721 | 0.038471 | 0.02 | 0.04 | 0.058109 | 0.12787 | 0.084853 | 0.15362 | 0.2393 | 0.1675 | 0.015055 | 0.095847 | 0.021093 | 0.072572 | 0.090835 | 0.023405 | 0.44098 | 0.03656 | 0.32462 | 0.027142 | 0.033912 | 0.041673 | 0.037639 | 0.056214 | 0.048442 | 0.02582
| Rel.Std.Dev. |  | 3.566% | 0.419% | 1.629% | 4.465% | 1.35% | 0.6993% | 1.351% | 1.955% | 4.51% | 3.009% | 5.428% | 8.971% | 5.915% | 0.524% | 3.419% | 0.7748% | 2.982% | 3.24% | 0.8118% | 16.68% | 1.198% | 13.98% | 0.9206% | 1.188% | 1.472% | 1.231% | 1.938% | 1.74% | 0.8347%
| PDM3 |  | 1.557% | -1.103% | -3.891% | 6.89% | -0.8771% | -0.5293% | 2.949% | 3.354% | -1.399% | -1.92% | -1.573% | -7.224% | -1.515% | -0.06557% | -2.5% | -5.318% | -15.37% | -2.483% | 0.2705% | -8.065% | 6.137% | -19.25% | 2.543% | -0.7032% | -1.515% | 6.369% | 0.8619% | -3.196% | 7.586%
| Z-Score (Absolute) |  | 0.71 | 0.07 | 0.88 | 1.74 | 0 | 0.1 | 1.11 | 1.23 | 0.15 | 0.3 | 0.2 | 1.15 | 0.19 | 0.24 | 0.47 | 1.29 | 4.21 | 0.47 | 0.33 | 1.85 | 2.04 | 6.47 | 0.99 | 0.05 | 0.34 | 2.11 | 0.3 | 0.67 | 2.46
| Analytical results for La in OREAS 905 (Certified Value 46 ppm) | [31 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 42.5 | 39.3 | 46.9 | 47.5 | 45.3 | 42.3 | 50.4 | 46 | 44.6 | 38.2 | 22 | 44.8 | 37 | 47.67 | 45 | 48.7192 | 48 | 48 | 42 | 49.1 | 49.4 | 45.2 | 42.4 | 44.9 | 48.1 | 47.9 |  |  |  |  | 1
| 1 | 2 | 43.5 | 40.8 | 49.1 | 47.6 | 45.3 | 44.2 | 50.5 | 54 | 41.9 | 36.5 | 24 | 45.9 | 38 | 48.81 | 44.9 | 48.7903 | 48 | 48 | 45 | 48.4 | 43.6 | 46.3 | 43.5 | 44.1 | 47.4 | 49.1 |  |  |  |  | 0
| 1 | 3 | 45.8 | 40.6 | 49.5 | 44.6 | 45.9 | 43.7 | 50.3 | 47 | 40.8 | 44.9 | < 20 | 43.6 | 37 | 47.32 | 44 | 48.3327 | 45 | 45 | 45 | 46.2 | 45.8 | 47 | 44 | 45.5 | 47.5 | 49.6 |  |  |  |  | 16
| 1 | 4 | 43.2 | 37.6 | 50.5 | 46.7 | 45.8 | 42.6 | 49.2 | 46 | 44.5 | 44.5 | 29 | 44.3 | 35 | 47.71 | 44.9 | 47.9897 | 45 | 47 | 45 | 48 | 44.1 | 45.5 | 44.5 | 44.9 | 48.1 | 49.7 |  |  |  |  | 45.9619
| 1 | 5 | 43.2 | 37.4 | 49.7 | 46.8 | 46.4 | 44.3 | 50.9 | 49 | 27.4 | 41.9 | 27 | 45.1 | 38 | 49.04 | 45.5 | 47.3883 | 45 | 49 | 46 | 47 | 49 | 46.2 | 43 | 46.2 | 47.4 | 49.8 |  |  |  |  | 32
| 1 | 6 | 43.6 | 38.2 | 49.9 | 46 | 45.8 | 45.4 | 51.1 | 50 | 37.4 | 47.3 | 35 | 45.6 | 36 | 48.39 | 43.8 | 47.1977 | 44 | 48 | 44 | 48.3 | 49.9 | 44.7 | 43.7 | 47.5 | 46.3 | 49.9
| Mean |  | 43.6333 | 38.9833 | 49.2667 | 46.5333 | 45.75 | 43.75 | 50.4 | 48.6667 | 39.4333 | 42.2167 | 27.4 | 44.8833 | 36.8333 | 48.1567 | 44.6833 | 48.0696 | 45.8333 | 47.5 | 44.5 | 47.8333 | 46.9667 | 45.8167 | 43.5167 | 45.5167 | 47.4667 | 49.3333
| Median |  | 43.35 | 38.75 | 49.6 | 46.75 | 45.8 | 43.95 | 50.45 | 48 | 41.35 | 43.2 | 27 | 44.95 | 37 | 48.05 | 44.9 | 48.1612 | 45 | 48 | 45 | 48.15 | 47.4 | 45.85 | 43.6 | 45.2 | 47.45 | 49.65
| Std Dev. |  | 1.129 | 1.4865 | 1.2485 | 1.113 | 0.41352 | 1.1537 | 0.66332 | 3.0768 | 6.4692 | 4.1754 | 5.0299 | 0.84715 | 1.169 | 0.69252 | 0.64936 | 0.66954 | 1.7224 | 1.3784 | 1.3784 | 1.052 | 2.8133 | 0.83766 | 0.7414 | 1.1974 | 0.65929 | 0.75542
| Rel.Std.Dev. |  | 2.587% | 3.813% | 2.534% | 2.392% | 0.9039% | 2.637% | 1.316% | 6.322% | 16.41% | 9.89% | 18.36% | 1.887% | 3.174% | 1.438% | 1.453% | 1.393% | 3.758% | 2.902% | 3.098% | 2.199% | 5.99% | 1.828% | 1.704% | 2.631% | 1.389% | 1.531%
| PDM3 |  | -5.066% | -15.18% | 7.19% | 1.243% | -0.4611% | -4.813% | 9.656% | 5.885% | -14.2% | -8.149% | -40.39% | -2.347% | -19.86% | 4.775% | -2.782% | 4.586% | -0.2798% | 3.346% | -3.181% | 4.072% | 2.186% | -0.316% | -5.32% | -0.9688% | 3.274% | 7.335%
| Z-Score (Absolute) |  | 0.81 | 2.12 | 1.24 | 0.23 | 0.01 | 0.64 | 1.44 | 0.9 | 1.23 | 0.76 | 5.74 | 0.28 | 2.8 | 0.74 | 0.34 | 0.71 | 0.02 | 0.54 | 0.24 | 0.64 | 0.37 | 0.01 | 0.71 | 0.08 | 0.53 | 1.2
| Analytical results for Li in OREAS 905 (Certified Value 20 ppm) | [31 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 1
2
| 1 | 1 | 21 | 20.1003 | 22 | 18 | 19.1 | 24.8 | 19.4 | 20 | 19 | 20 | 20 | 19.7 | 23 | 19 | 19 | 19.4 | 17 | 19.4 | 18.408 | 24.2 | 22.5 | 17.9 | 21.2 | 20 | 19.1 | 19 | 21.1 |  |  |  | 1
| 1 | 2 | 22.5 | 20.2546 | 23 | 19 | 19.2 | 25.2 | 20.2 | 20 | 19 | 22 | 19 | 23 | 23 | 20 | 19 | 19.7 | 17 | 18.1 | 19.581 | 26.8 | 21.5 | 16.6 | 22.7 | 19 | 18.6 | 19 | 20.7 |  |  |  | 0
| 1 | 3 | 23.7 | 20.2024 | 24 | 19 | 18.7 | 24.9 | 20 | 20 | 19 | 20 | 20 | 21.2 | 23 | 19 | 19 | 18.8 | 16 | 18.1 | 20.146 | 28.2 | 22 | 20.6 | 23 | 20 | 18.8 | 19 | 21.7 |  |  |  | 16
| 1 | 4 | 22.3 | 20.3204 | 23 | 19 | 19.2 | 24.8 | 20.1 | 20 | 19 | 20 | 20 | 21.3 | 22 | 20 | 18 | 19.3 | 17 | 19.3 | 20.353 | 26.8 | 21 | 20.9 | 22.8 | 20 | 19.1 | 19 | 20.4 |  |  |  | 20.0438
| 1 | 5 | 22 | 20.384 | 24 | 17 | 19.3 | 24.9 | 20 | 20 | 20 | 20 | 20 | 20.5 | 23 | 20 | 18 | 18.3 | 17 | 19.9 | 19.561 | 26.3 | 20.5 | 18.6 | 23.1 | 20 | 18.9 | 19 | 22 |  |  |  | 33
| 1 | 6 | 21.4 | 20.393 | 22 | 19 | 18.9 | 24.5 | 19.8 | 20 | 20 | 20 | 20 | 20.7 | 23 | 20 | 18 | 19.1 | 18 | 19.2 | 19.026 | 19.6 | 21.5 | 19.4 | 22 | 18 | 19.7 | 19 | 21
| Mean |  | 22.15 | 20.2758 | 23 | 18.5 | 19.0667 | 24.85 | 19.9167 | 20 | 19.3333 | 20.3333 | 19.8333 | 21.0667 | 22.8333 | 19.6667 | 18.5 | 19.1 | 17 | 19 | 19.5125 | 25.3167 | 21.5 | 19 | 22.4667 | 19.5 | 19.0333 | 19 | 21.15
| Median |  | 22.15 | 20.2875 | 23 | 19 | 19.15 | 24.85 | 20 | 20 | 19 | 20 | 20 | 20.95 | 23 | 20 | 18.5 | 19.2 | 17 | 19.25 | 19.571 | 26.55 | 21.5 | 19 | 22.75 | 20 | 19 | 19 | 21.05
| Std Dev. |  | 0.94393 | 0.11322 | 0.89443 | 0.83666 | 0.22509 | 0.22583 | 0.28577 | 0 | 0.5164 | 0.8165 | 0.40825 | 1.1075 | 0.40825 | 0.5164 | 0.54772 | 0.49396 | 0.63246 | 0.73756 | 0.71713 | 3.0857 | 0.70711 | 1.6407 | 0.73121 | 0.83666 | 0.37771 | 0 | 0.60249
| Rel.Std.Dev. |  | 4.262% | 0.5584% | 3.889% | 4.522% | 1.181% | 0.9088% | 1.435% | 0% | 2.671% | 4.016% | 2.058% | 5.257% | 1.788% | 2.626% | 2.961% | 2.586% | 3.72% | 3.882% | 3.675% | 12.19% | 3.289% | 8.635% | 3.255% | 4.291% | 1.984% | 0% | 2.849%
| PDM3 |  | 10.51% | 1.157% | 14.75% | -7.702% | -4.875% | 23.98% | -0.6343% | -0.2185% | -3.545% | 1.445% | -1.05% | 5.103% | 13.92% | -1.882% | -7.702% | -4.709% | -15.19% | -5.208% | -2.651% | 26.31% | 7.265% | -5.208% | 12.09% | -2.713% | -5.041% | -5.208% | 5.519%
| Z-Score (Absolute) |  | 1.87 | 0.36 | 2.56 | 1.08 | 0.62 | 4.06 | 0.15 | 0.13 | 0.4 | 0.4 | 0 | 1 | 2.43 | 0.13 | 1.08 | 0.59 | 2.29 | 0.67 | 0.26 | 5.36 | 1.35 | 0.67 | 2.34 | 0.27 | 0.65 | 0.67 | 1.07
| Analytical results for Lu in OREAS 905 (Certified Value 0.1 ppm) | [31 empty] | Certified
Lu | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  |  | 1
| No | No | 01: 04; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 08: 15; 09: 17; 10: 26; 11: 30; 12: 32; 13: 33 |  |  |  | ppm
|  |  | 01: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*OES/MS |  |  |  | 2
2
| 1 | 1 | 01: 0.09; 03: 0.09; 04: 0.1; 05: 0.1; 06: < 0.1; 07: < 0.1; 08: 0.1; 09: 0.09; 10: 0.1; 11: 0.12; 12: 0.11; 13: < 0.1 |  |  |  | 1
| 1 | 2 | 01: 0.09; 03: 0.1; 04: 0.18; 05: 0.1; 06: < 0.1; 07: < 0.1; 08: 0.12; 09: 0.08; 10: 0.08; 11: 0.13; 12: 0.1; 13: < 0.1 |  |  |  | 6
| 1 | 3 | 01: 0.08; 03: 0.1; 04: 0.1; 05: 0.1; 06: < 0.1; 07: < 0.1; 08: 0.12; 09: 0.08; 10: 0.08; 11: 0.13; 12: 0.11; 13: < 0.1 |  |  |  | 16
| 1 | 4 | 01: 0.09; 03: 0.1; 04: 0.1; 05: 0.1; 06: < 0.1; 07: < 0.1; 08: 0.12; 09: 0.09; 10: 0.08; 11: 0.14; 12: 0.11; 13: < 0.1 |  |  |  | 0.10111
| 1 | 5 | 01: 0.09; 03: 0.1; 04: 0.1; 05: 0.1; 06: < 0.1; 07: < 0.1; 08: 0.12; 09: 0.09; 10: 0.1; 11: 0.12; 12: 0.09; 13: < 0.1 |  |  |  | 34
| 1 | 6 | 01: 0.08; 03: 0.1; 04: 0.1; 05: 0.1; 06: < 0.1; 07: < 0.1; 08: 0.13; 09: 0.09; 10: 0.1; 11: 0.12; 12: 0.1; 13: < 0.1
| Mean |  | 01: 0.086667; 03: 0.098333; 04: 0.11333; 05: 0.1; 08: 0.11833; 09: 0.086667; 10: 0.09; 11: 0.12667; 12: 0.10333
| Median |  | 01: 0.09; 03: 0.1; 04: 0.1; 05: 0.1; 08: 0.12; 09: 0.09; 10: 0.09; 11: 0.125; 12: 0.105
| Std Dev. |  | 01: 0.005164; 03: 0.0040825; 04: 0.03266; 05: 1.5202e-17; 08: 0.0098319; 09: 0.005164; 10: 0.010954; 11: 0.008165; 12: 0.008165
| Rel.Std.Dev. |  | 01: 5.958%; 03: 4.152%; 04: 28.82%; 05: 1.52e-14%; 08: 8.309%; 09: 5.958%; 10: 12.17%; 11: 6.446%; 12: 7.902%
| PDM3 |  | 01: -14.29%; 03: -2.747%; 04: 12.09%; 05: -1.099%; 08: 17.03%; 09: -14.29%; 10: -10.99%; 11: 25.27%; 12: 2.198%
| Z-Score (Absolute) |  | 01: 0.9; 03: 0.11; 04: 0; 05: 0; 06: NA; 07: NA; 08: 1.24; 09: 0.9; 10: 0.67; 11: 1.8; 12: 0.22; 13: NA
| Analytical results for Mg in OREAS 905 (Certified Value 0.276 wt.%) | [31 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.25 | 0.28259 | 0.28 | 0.3 | 0.27 | 0.27 | 0.281 | 0.29 | 0.27 | 0.28 | 0.27 | 0.2542 | 0.25 | 0.27 | 0.25 | 0.2653 | 0.27 | 0.29 | 0.28449 | 0.3 | 0.29 | 0.29 | 0.26 | 0.26 | 0.292 | 0.26 | 0.27 | 0.3 |  |  | 1
| 1 | 2 | 0.25 | 0.28961 | 0.29 | 0.311 | 0.26 | 0.27 | 0.287 | 0.28 | 0.27 | 0.3 | 0.27 | 0.2451 | 0.26 | 0.27 | 0.26 | 0.265 | 0.28 | 0.28 | 0.28462 | 0.3 | 0.29 | 0.3 | 0.26 | 0.26 | 0.293 | 0.26 | 0.27 | 0.29 |  |  | 0
| 1 | 3 | 0.27 | 0.28138 | 0.3 | 0.302 | 0.26 | 0.27 | 0.284 | 0.29 | 0.27 | 0.28 | 0.28 | 0.2852 | 0.24 | 0.27 | 0.25 | 0.2631 | 0.23 | 0.28 | 0.28535 | 0.29 | 0.28 | 0.3 | 0.26 | 0.26 | 0.295 | 0.26 | 0.26 | 0.3 |  |  | 16
| 1 | 4 | 0.27 | 0.28688 | 0.29 | 0.304 | 0.26 | 0.27 | 0.282 | 0.28 | 0.26 | 0.27 | 0.28 | 0.2811 | 0.26 | 0.27 | 0.25 | 0.2605 | 0.27 | 0.28 | 0.28154 | 0.29 | 0.28 | 0.3 | 0.25 | 0.26 | 0.298 | 0.26 | 0.27 | 0.3 |  |  | 0.27622
| 1 | 5 | 0.26 | 0.28512 | 0.3 | 0.303 | 0.26 | 0.28 | 0.291 | 0.3 | 0.27 | 0.29 | 0.26 | 0.2679 | 0.26 | 0.27 | 0.25 | 0.268 | 0.27 | 0.29 | 0.28056 | 0.28 | 0.29 | 0.3 | 0.26 | 0.25 | 0.292 | 0.26 | 0.27 | 0.3 |  |  | 35
| 1 | 6 | 0.26 | 0.2895 | 0.28 | 0.316 | 0.27 | 0.28 | 0.281 | 0.29 | 0.27 | 0.3 | 0.27 | 0.2839 | 0.27 | 0.26 | 0.24 | 0.2664 | 0.28 | 0.28 | 0.28136 | 0.28 | 0.27 | 0.29 | 0.26 | 0.25 | 0.29 | 0.27 | 0.27 | 0.3
| Mean |  | 0.26 | 0.28585 | 0.29 | 0.306 | 0.26333 | 0.27333 | 0.28433 | 0.28833 | 0.26833 | 0.28667 | 0.27167 | 0.26957 | 0.25667 | 0.26833 | 0.25 | 0.26472 | 0.26667 | 0.28333 | 0.28299 | 0.29 | 0.28333 | 0.29667 | 0.25833 | 0.25667 | 0.29333 | 0.26167 | 0.26833 | 0.29833
| Median |  | 0.26 | 0.286 | 0.29 | 0.3035 | 0.26 | 0.27 | 0.283 | 0.29 | 0.27 | 0.285 | 0.27 | 0.2745 | 0.26 | 0.27 | 0.25 | 0.26515 | 0.27 | 0.28 | 0.28301 | 0.29 | 0.285 | 0.3 | 0.26 | 0.26 | 0.2925 | 0.26 | 0.27 | 0.3
| Std Dev. |  | 0.0089443 | 0.0034566 | 0.0089443 | 0.0061644 | 0.005164 | 0.005164 | 0.0039833 | 0.0075277 | 0.0040825 | 0.012111 | 0.0075277 | 0.016855 | 0.010328 | 0.0040825 | 0.0063246 | 0.0026241 | 0.018619 | 0.005164 | 0.0020566 | 0.0089443 | 0.008165 | 0.005164 | 0.0040825 | 0.005164 | 0.0028048 | 0.0040825 | 0.0040825 | 0.0040825
| Rel.Std.Dev. |  | 3.44% | 1.209% | 3.084% | 2.015% | 1.961% | 1.889% | 1.401% | 2.611% | 1.521% | 4.225% | 2.771% | 6.252% | 4.024% | 1.521% | 2.53% | 0.9913% | 6.982% | 1.823% | 0.7268% | 3.084% | 2.882% | 1.741% | 1.58% | 2.012% | 0.9562% | 1.56% | 1.521% | 1.368%
| PDM3 |  | -5.872% | 3.485% | 4.989% | 10.78% | -4.665% | -1.045% | 2.938% | 4.386% | -2.855% | 3.783% | -1.648% | -2.408% | -7.078% | -2.855% | -9.492% | -4.164% | -3.458% | 2.576% | 2.451% | 4.989% | 2.576% | 7.403% | -6.475% | -7.078% | 6.196% | -5.268% | -2.855% | 8.006%
| Z-Score (Absolute) |  | 0.76 | 0.68 | 0.91 | 1.8 | 0.58 | 0.02 | 0.59 | 0.82 | 0.3 | 0.73 | 0.11 | 0.23 | 0.95 | 0.3 | 1.32 | 0.5 | 0.02 | 0.54 | 0.52 | 0.91 | 0.54 | 1.28 | 0.86 | 0.95 | 1.1 | 0.67 | 0.3 | 1.38
| Analytical results for Mn in OREAS 905 (Certified Value 0.038 wt.%) | [31 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 3
3
| 1 | 1 | 0.0358 | 0.040519 | 0.0378 | 0.0392 | 0.0377 | 0.0369 | 0.0368 | 0.0392 | 0.0381 | 0.0363 | 0.039 | 0.0378 | 0.04 | 0.0358 | 0.0368 | 0.0385 | 0.0344 | 0.0406 | 0.038501 | 0.0409 | 0.0341 | 0.0379 | 0.0374 | 0.0366 | 0.0367 | 0.0393 | 0.0366 | 0.0365 | 0.0391 |  | 1
| 1 | 2 | 0.0369 | 0.040017 | 0.0388 | 0.0399 | 0.0374 | 0.0373 | 0.0375 | 0.0386 | 0.037 | 0.0396 | 0.0393 | 0.0373 | 0.04 | 0.0367 | 0.0376 | 0.0375 | 0.0357 | 0.0404 | 0.038229 | 0.0404 | 0.0347 | 0.0382 | 0.0382 | 0.0369 | 0.0371 | 0.0389 | 0.0366 | 0.0357 | 0.0387 |  | 0
| 1 | 3 | 0.0398 | 0.039581 | 0.0391 | 0.039 | 0.0361 | 0.0377 | 0.0371 | 0.0392 | 0.037 | 0.0375 | 0.0396 | 0.039 | 0.04 | 0.0354 | 0.0364 | 0.0367 | 0.032 | 0.0393 | 0.037729 | 0.0388 | 0.0333 | 0.0384 | 0.0376 | 0.0369 | 0.0375 | 0.0395 | 0.0372 | 0.0362 | 0.0395 |  | 16
| 1 | 4 | 0.0394 | 0.040077 | 0.0388 | 0.0396 | 0.0371 | 0.0373 | 0.0368 | 0.0392 | 0.036 | 0.0369 | 0.0402 | 0.0367 | 0.04 | 0.0358 | 0.0364 | 0.0369 | 0.0349 | 0.0406 | 0.037827 | 0.0383 | 0.0345 | 0.038 | 0.0372 | 0.0359 | 0.0373 | 0.0399 | 0.0362 | 0.0363 | 0.0405 |  | 0.037747
| 1 | 5 | 0.0381 | 0.039549 | 0.0385 | 0.0391 | 0.0373 | 0.0391 | 0.0379 | 0.0406 | 0.0373 | 0.0371 | 0.0393 | 0.0355 | 0.04 | 0.0359 | 0.0372 | 0.0376 | 0.0343 | 0.04 | 0.038158 | 0.0385 | 0.0355 | 0.0382 | 0.037 | 0.037 | 0.0365 | 0.0385 | 0.0369 | 0.0355 | 0.0393 |  | 36
| 1 | 6 | 0.039 | 0.04027 | 0.0377 | 0.0406 | 0.0383 | 0.0377 | 0.0368 | 0.0406 | 0.037 | 0.0374 | 0.0398 | 0.0365 | 0.04 | 0.0355 | 0.0352 | 0.0371 | 0.0335 | 0.0403 | 0.037745 | 0.0369 | 0.0348 | 0.0382 | 0.0372 | 0.037 | 0.0362 | 0.0383 | 0.0377 | 0.0355 | 0.0391
| Mean |  | 0.038167 | 0.040002 | 0.03845 | 0.039567 | 0.037317 | 0.037667 | 0.03715 | 0.039567 | 0.037067 | 0.037467 | 0.039533 | 0.037133 | 0.04 | 0.03585 | 0.0366 | 0.037383 | 0.034133 | 0.0402 | 0.038031 | 0.038967 | 0.034483 | 0.03815 | 0.037433 | 0.036717 | 0.036883 | 0.039067 | 0.036867 | 0.03595 | 0.039367
| Median |  | 0.03855 | 0.040047 | 0.03865 | 0.0394 | 0.03735 | 0.0375 | 0.03695 | 0.0392 | 0.037 | 0.03725 | 0.03945 | 0.037 | 0.04 | 0.0358 | 0.0366 | 0.0373 | 0.03435 | 0.04035 | 0.037992 | 0.03865 | 0.0346 | 0.0382 | 0.0373 | 0.0369 | 0.0369 | 0.0391 | 0.03675 | 0.03595 | 0.0392
| Std Dev. |  | 0.0015552 | 0.00038119 | 0.00057533 | 0.00060882 | 0.00072778 | 0.00076333 | 0.00045935 | 0.00083347 | 0.00067429 | 0.001129 | 0.0004274 | 0.0012011 | 0 | 0.00045935 | 0.00082946 | 0.00064627 | 0.0012723 | 0.00049396 | 0.00031333 | 0.0014665 | 0.00073869 | 0.00017607 | 0.0004274 | 0.00042622 | 0.00049967 | 0.0006121 | 0.00052789 | 0.00043704 | 0.00061536
| Rel.Std.Dev. |  | 4.075% | 0.9529% | 1.496% | 1.539% | 1.95% | 2.027% | 1.236% | 2.106% | 1.819% | 3.013% | 1.081% | 3.235% | 0% | 1.281% | 2.266% | 1.729% | 3.727% | 1.229% | 0.8239% | 3.764% | 2.142% | 0.4615% | 1.142% | 1.161% | 1.355% | 1.567% | 1.432% | 1.216% | 1.563%
| PDM3 |  | 1.113% | 5.975% | 1.863% | 4.822% | -1.139% | -0.2119% | -1.581% | 4.822% | -1.801% | -0.7418% | 4.733% | -1.625% | 5.97% | -5.025% | -3.038% | -0.9626% | -9.573% | 6.499% | 0.7542% | 3.232% | -8.645% | 1.069% | -0.8301% | -2.729% | -2.287% | 3.497% | -2.331% | -4.76% | 4.292%
| Z-Score (Absolute) |  | 0.67 | 2.25 | 0.92 | 1.88 | 0.06 | 0 | 0.2 | 1.88 | 0.27 | 0.3 | 1.85 | 0.22 | 2.25 | 1.32 | 0.67 | 0 | 2.43 | 2.42 | 0.56 | 1.36 | 2.5 | 0.66 | 0.04 | 0.57 | 0.43 | 1.45 | 0.44 | 1.23 | 1.51
| Analytical results for Mo in OREAS 905 (Certified Value 3.27 ppm) | [31 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 3.05 | 3.2984 | 3.53 | 3.4 | 3.27 | 3.12 | 3.2 | 3 | 3.08 | 3 | 3.95 | 3.1 | 4 | 3.42 | 3 | 3.5 | 10 | 4.7 | 2.8618 | < 2 | 3.7 | 2 | 4 | 3.19 | 3.22 | 3.6 | 3.28 | 3.54 | 3.55 |  | 1
| 1 | 2 | 3.18 | 3.2731 | 3.97 | 3.2 | 3.22 | 3.22 | 3.2 | 3 | 4.16 | 3 | 3.05 | 3.1 | 4 | 3.39 | 3 | 3.4 | 6 | 3.7 | 3.0364 | < 2 | 3.4 | 2 | 4 | 2.7 | 3.2 | 3.5 | 3.08 | 3.42 | 3.15 |  | 7
| 1 | 3 | 3.19 | 3.3912 | 3.64 | 3.3 | 2.92 | 3.22 | 3.2 | 3 | 2.91 | 3 | 3.41 | 3.5 | 4 | 3.37 | 4 | 3.2 | 7 | 3.2 | 3.0464 | < 2 | 2.9 | 3 | 3.5 | 3.4 | 3.35 | 3.8 | 3.17 | 3.2 | 3.62 |  | 16
| 1 | 4 | 3.05 | 3.2099 | 3.67 | 3.4 | 3 | 3.21 | 3.3 | 3 | 2.69 | 3 | 3.17 | 3.1 | < 3 | 3.33 | < 2 | 3.3 | 7 | 3.4 | 3.1273 | 2 | 3.4 | < 2 | 3.5 | 3.44 | 3.22 | 3.6 | 3.18 | 3.5 | 3.47 |  | 3.2703
| 1 | 5 | 3.09 | 3.2295 | 3.97 | 3.3 | 3 | 3.24 | 3.3 | 3 | 3.05 | 3 | 3.43 | 3 | 5 | 3.35 | 2 | 3.3 | 7 | 3.2 | 2.7155 | < 2 | 3.4 | 2 | 3.5 | 3.1 | 3.26 | 3.6 | 3.27 | 3.46 | 3.44 |  | 37
| 1 | 6 | 3.17 | 3.3574 | 3.45 | 3.3 | 3.04 | 3.26 | 3.1 | 3 | 3.05 | 3 | 2.95 | 3.5 | 5 | 3.47 | 3 | 3.2 | 8 | 3.3 | 2.7382 | < 2 | 3.2 | < 2 | 3.5 | 3.27 | 3.19 | 3.6 | 3.27 | 3.36 | 3.32
| Mean |  | 3.1217 | 3.2932 | 3.705 | 3.3167 | 3.075 | 3.2117 | 3.2167 | 3 | 3.1567 | 3 | 3.3267 | 3.2167 | 4.4 | 3.3883 | 3 | 3.3167 | 7.5 | 3.5833 | 2.9209 | 2 | 3.3333 | 2.25 | 3.6667 | 3.1833 | 3.24 | 3.6167 | 3.2083 | 3.4133 | 3.425
| Median |  | 3.13 | 3.2858 | 3.655 | 3.3 | 3.02 | 3.22 | 3.2 | 3 | 3.05 | 3 | 3.29 | 3.1 | 4 | 3.38 | 3 | 3.3 | 7 | 3.35 | 2.9491 | 2 | 3.4 | 2 | 3.5 | 3.23 | 3.22 | 3.6 | 3.225 | 3.44 | 3.455
| Std Dev. |  | 0.065853 | 0.070914 | 0.2198 | 0.075277 | 0.13824 | 0.048339 | 0.075277 | 0 | 0.51263 | 0 | 0.36031 | 0.22286 | 0.54772 | 0.050761 | 0.70711 | 0.1169 | 1.3784 | 0.57764 | 0.17364 |  | 0.26583 | 0.5 | 0.2582 | 0.26868 | 0.058992 | 0.098319 | 0.079352 | 0.12176 | 0.16885
| Rel.Std.Dev. |  | 2.11% | 2.153% | 5.932% | 2.27% | 4.496% | 1.505% | 2.34% | 0% | 16.24% | 0% | 10.83% | 6.928% | 12.45% | 1.498% | 23.57% | 3.525% | 18.38% | 16.12% | 5.945% |  | 7.975% | 22.22% | 7.042% | 8.44% | 1.821% | 2.719% | 2.473% | 3.567% | 4.93%
| PDM3 |  | -4.546% | 0.7006% | 13.29% | 1.417% | -5.973% | -1.794% | -1.641% | -8.266% | -3.475% | -8.266% | 1.723% | -1.641% | 34.54% | 3.608% | -8.266% | 1.417% | 129.3% | 9.571% | -10.68% | -38.84% | 1.927% | -31.2% | 12.12% | -2.66% | -0.9273% | 10.59% | -1.896% | 4.373% | 4.73%
| Z-Score (Absolute) |  | 0.63 | 0.21 | 2.22 | 0.32 | 0.85 | 0.1 | 0.16 | 1.22 | 1.43 | 1.22 | 0.37 | 0.16 | 5.61 | 0.67 | 0 | 0.32 | 18.28 | 0.54 | 1.6 | 6.09 | 0.41 | 4.87 | 2.03 | 0.32 | 0.16 | 1.79 | 0.2 | 0.8 | 0.85
| Analytical results for Na in OREAS 905 (Certified Value 2.4 wt.%) | [31 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 2.25 | 2.6848 | 2.49 | 2.32 | 2.38 | 2.46 | 2.56 | 2.41 | 2.32 | 2.37 | 2.36 | 2.2606 | 2.74 | 2.24 | 2.23 | 2.3513 | 2.27 | 2.515 | 2.4163 | 2.51 | 2.39 | 2.44 | 2.61 | 2.39 | 2.27 | 2.57 | 2.3 | 2.4 | 2.512 |  | 1
| 1 | 2 | 2.29 | 2.6633 | 2.53 | 2.37 | 2.38 | 2.46 | 2.55 | 2.4 | 2.31 | 2.48 | 2.35 | 2.2939 | 2.85 | 2.23 | 2.26 | 2.3753 | 2.17 | 2.494 | 2.4317 | 2.54 | 2.54 | 2.42 | 2.62 | 2.39 | 2.33 | 2.58 | 2.25 | 2.4 | 2.498 |  | 8
| 1 | 3 | 2.46 | 2.6646 | 2.46 | 2.33 | 2.33 | 2.47 | 2.53 | 2.4 | 2.31 | 2.33 | 2.4 | 2.3937 | 2.69 | 2.25 | 2.25 | 2.3523 | 2.12 | 2.505 | 2.3833 | 2.52 | 2.41 | 2.52 | 2.59 | 2.36 | 2.38 | 2.62 | 2.29 | 2.32 | 2.523 |  | 16
| 1 | 4 | 2.4 | 2.7212 | 2.46 | 2.37 | 2.36 | 2.45 | 2.52 | 2.38 | 2.23 | 2.3 | 2.39 | 2.3052 | 2.72 | 2.23 | 2.16 | 2.3416 | 2.17 | 2.612 | 2.4118 | 2.54 | 2.4 | 2.45 | 2.59 | 2.34 | 2.33 | 2.65 | 2.33 | 2.44 | 2.532 |  | 2.3997
| 1 | 5 | 2.37 | 2.7105 | 2.49 | 2.45 | 2.37 | 2.51 | 2.57 | 2.48 | 2.35 | 2.39 | 2.41 | 2.1703 | 2.8 | 2.26 | 2.17 | 2.3794 | 2.23 | 2.462 | 2.3673 | 2.52 | 2.47 | 2.46 | 2.61 | 2.39 | 2.29 | 2.58 | 2.28 | 2.33 | 2.563 |  | 38
| 1 | 6 | 2.37 | 2.7297 | 2.42 | 2.4 | 2.42 | 2.49 | 2.52 | 2.44 | 2.38 | 2.39 | 2.4 | 2.4877 | 2.74 | 2.23 | 2.12 | 2.4059 | 2.29 | 2.483 | 2.3969 | 2.48 | 2.36 | 2.42 | 2.6 | 2.35 | 2.27 | 2.56 | 2.35 | 2.35 | 2.524
| Mean |  | 2.3567 | 2.6957 | 2.475 | 2.3733 | 2.3733 | 2.4733 | 2.5417 | 2.4183 | 2.3167 | 2.3767 | 2.385 | 2.3186 | 2.7567 | 2.24 | 2.1983 | 2.3676 | 2.2083 | 2.5118 | 2.4012 | 2.5183 | 2.4283 | 2.4517 | 2.6033 | 2.37 | 2.3117 | 2.5933 | 2.3 | 2.3733 | 2.5253
| Median |  | 2.37 | 2.6977 | 2.475 | 2.37 | 2.375 | 2.465 | 2.54 | 2.405 | 2.315 | 2.38 | 2.395 | 2.2995 | 2.74 | 2.235 | 2.2 | 2.3638 | 2.2 | 2.4995 | 2.4043 | 2.52 | 2.405 | 2.445 | 2.605 | 2.375 | 2.31 | 2.58 | 2.295 | 2.375 | 2.5235
| Std Dev. |  | 0.075807 | 0.028827 | 0.037283 | 0.04761 | 0.029439 | 0.022509 | 0.02137 | 0.036009 | 0.050465 | 0.06186 | 0.02429 | 0.10989 | 0.058195 | 0.012649 | 0.056362 | 0.023838 | 0.065853 | 0.052389 | 0.023489 | 0.022286 | 0.065549 | 0.037103 | 0.012111 | 0.022804 | 0.043089 | 0.034448 | 0.035777 | 0.047188 | 0.021888
| Rel.Std.Dev. |  | 3.217% | 1.069% | 1.506% | 2.006% | 1.24% | 0.9101% | 0.8408% | 1.489% | 2.178% | 2.603% | 1.018% | 4.74% | 2.111% | 0.5647% | 2.564% | 1.007% | 2.982% | 2.086% | 0.9782% | 0.885% | 2.699% | 1.513% | 0.4652% | 0.9622% | 1.864% | 1.328% | 1.556% | 1.988% | 0.8667%
| PDM3 |  | -1.792% | 12.34% | 3.139% | -1.098% | -1.098% | 3.07% | 5.917% | 0.7776% | -3.459% | -0.9588% | -0.6115% | -3.38% | 14.88% | -6.654% | -8.39% | -1.335% | -7.974% | 4.674% | 0.06378% | 4.945% | 1.194% | 2.167% | 8.487% | -1.237% | -3.667% | 8.07% | -4.154% | -1.098% | 5.237%
| Z-Score (Absolute) |  | 0.26 | 2.86 | 0.83 | 0.11 | 0.11 | 0.81 | 1.44 | 0.31 | 0.63 | 0.08 | 0 | 0.61 | 3.42 | 1.33 | 1.72 | 0.16 | 1.62 | 0.98 | 0.15 | 1.23 | 0.4 | 0.61 | 2.01 | 0.14 | 0.67 | 1.92 | 0.78 | 0.11 | 1.29
| Analytical results for Nb in OREAS 905 (Certified Value 18.1 ppm) | [31 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 17.3 | 18.5712 | 16.3 | 20.5 | 18.4 | 19.3 | 18.9 | 19 | 20 | 19 | 12.9 | 17.1 | 19.5 | 34 | 17.92 | 8.5 | 17 | 18.4 | 15 | 18.5 | 17.1 | 18.3 | 14.8 | 17.8 | 20.2 | 17.03 |  |  |  |  | 1
| 1 | 2 | 17.7 | 18.5324 | 18.3 | 19.5 | 18.2 | 19.3 | 18.3 | 19.5 | 31.4 | 19 | 16.2 | 17.7 | 19.4 | 33 | 18.14 | 8.5 | 17 | 17.7 | 14 | 17.5 | 15.2 | 18.9 | 11.9 | 18.1 | 20.1 | 17.89 |  |  |  |  | 15
| 1 | 3 | 18.3 | 18.5202 | 17.7 | 20.6 | 17.6 | 19.3 | 19.2 | 19 | 16.6 | 17 | 18 | 18 | 19 | 33 | 17.36 | 8.5 | 16 | 18.1 | 16 | 17.5 | 17.1 | 19.3 | 13.8 | 18 | 19.7 | 17.96 |  |  |  |  | 16
| 1 | 4 | 17.4 | 18.6704 | 18.4 | 19.6 | 18 | 19.5 | 17.1 | 19 | 15.5 | 18 | 16.6 | 17.4 | 19.4 | 32 | 16.92 | 8.7 | 16 | 17.3 | 15 | 18 | 17.7 | 19.1 | 12.7 | 17.8 | 19.3 | 17.59 |  |  |  |  | 18.1385
| 1 | 5 | 17.6 | 18.3336 | 17.6 | 19.2 | 17.9 | 19.8 | 18.7 | 20 | 18.2 | 17 | 16.6 | 16.6 | 19.7 | 32 | 17.87 | 7.6 | 17 | 18.1 | 15 | 17.5 | 17.2 | 19 | 11.5 | 18.6 | 20.3 | 17.74 |  |  |  |  | 39
| 1 | 6 | 17.5 | 18.6806 | 17.8 | 19.8 | 17.9 | 19.4 | 17.8 | 19 | 15.7 | 19 | 16.5 | 17.4 | 19.7 | 32 | 17.58 | 9.3 | 16 | 18 | 16 | 17.5 | 17.3 | 18.2 | 11.1 | 19.1 | 19.6 | 17.52
| Mean |  | 17.6333 | 18.5514 | 17.6833 | 19.8667 | 18 | 19.4333 | 18.3333 | 19.25 | 19.5667 | 18.1667 | 16.1333 | 17.3667 | 19.45 | 32.6667 | 17.6317 | 8.5167 | 16.5 | 17.9333 | 15.1667 | 17.75 | 16.9333 | 18.8 | 12.6333 | 18.2333 | 19.8667 | 17.6217
| Median |  | 17.55 | 18.5518 | 17.75 | 19.7 | 17.95 | 19.35 | 18.5 | 19 | 17.4 | 18.5 | 16.55 | 17.4 | 19.45 | 32.5 | 17.725 | 8.5 | 16.5 | 18.05 | 15 | 17.5 | 17.15 | 18.95 | 12.3 | 18.05 | 19.9 | 17.665
| Std Dev. |  | 0.3559 | 0.12651 | 0.75211 | 0.56451 | 0.27568 | 0.19664 | 0.77632 | 0.41833 | 6.0394 | 0.98319 | 1.7037 | 0.48442 | 0.25884 | 0.8165 | 0.44301 | 0.54559 | 0.54772 | 0.38297 | 0.75277 | 0.41833 | 0.87788 | 0.44721 | 1.4306 | 0.5164 | 0.39328 | 0.33523
| Rel.Std.Dev. |  | 2.018% | 0.682% | 4.253% | 2.841% | 1.532% | 1.012% | 4.234% | 2.173% | 30.87% | 5.412% | 10.56% | 2.789% | 1.331% | 2.499% | 2.513% | 6.406% | 3.32% | 2.136% | 4.963% | 2.357% | 5.184% | 2.379% | 11.32% | 2.832% | 1.98% | 1.902%
| PDM3 |  | -2.785% | 2.276% | -2.51% | 9.527% | -0.7638% | 7.138% | 1.074% | 6.128% | 7.873% | 0.155% | -11.05% | -4.255% | 7.23% | 80.1% | -2.795% | -53.05% | -9.034% | -1.131% | -16.38% | -2.142% | -6.644% | 3.647% | -30.35% | 0.5226% | 9.527% | -2.85%
| Z-Score (Absolute) |  | 0.31 | 0.74 | 0.12 | 2.05 | 0.19 | 1.62 | 0.52 | 1.44 | 0.61 | 0.36 | 1.03 | 0.44 | 1.64 | 14.83 | 0.18 | 9.27 | 1.31 | 0.12 | 2.64 | 0.21 | 0.53 | 0.99 | 5.16 | 0.42 | 2.05 | 0.19
| Analytical results for Nd in OREAS 905 (Certified Value 39.1 ppm) | [31 empty] | Certified
Nd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  |  |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 17; 08: 20; 09: 26; 10: 30; 11: 33 |  |  |  |  | ppm
|  |  | 01: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 01: 38.7; 03: 35.1; 04: 38.7; 05: 35.8; 06: 35.8; 07: 37.3; 08: 42.3188; 09: 39.8; 10: 38.3; 11: 40.4 |  |  |  |  | 1
| 1 | 2 | 01: 39.8; 03: 40; 04: 38.5; 05: 33.7; 06: 35; 07: 38.48; 08: 42.7716; 09: 38.4; 10: 39; 11: 39.6 |  |  |  |  | 0
| 1 | 3 | 01: 39.7; 03: 40.8; 04: 38; 05: 32.2; 06: 40.8; 07: 37.45; 08: 42.4028; 09: 39.2; 10: 38.8; 11: 41.2 |  |  |  |  | 16
| 1 | 4 | 01: 40.6; 03: 38.5; 04: 36.9; 05: 34.4; 06: 39.8; 07: 37.73; 08: 43.3941; 09: 39.3; 10: 41.4; 11: 39.4 |  |  |  |  | 39.0542
| 1 | 5 | 01: 40.1; 03: 38.6; 04: 38.3; 05: 23.8; 06: 39.3; 07: 38.29; 08: 42.5513; 09: 39.5; 10: 38.9; 11: 41.8 |  |  |  |  | 40
| 1 | 6 | 01: 40.3; 03: 41.9; 04: 38.5; 05: 30.5; 06: 41.3; 07: 37.79; 08: 42.5154; 09: 38.7; 10: 39.8; 11: 41.5
| Mean |  | 01: 39.8667; 03: 39.15; 04: 38.15; 05: 31.7333; 06: 38.6667; 07: 37.84; 08: 42.659; 09: 39.15; 10: 39.3667; 11: 40.65
| Median |  | 01: 39.95; 03: 39.3; 04: 38.4; 05: 32.95; 06: 39.55; 07: 37.76; 08: 42.5333; 09: 39.25; 10: 38.95; 11: 40.8
| Std Dev. |  | 01: 0.65929; 03: 2.3721; 04: 0.65651; 05: 4.294; 06: 2.6394; 07: 0.46269; 08: 0.39156; 09: 0.51672; 10: 1.1075; 11: 1.0075
| Rel.Std.Dev. |  | 01: 1.654%; 03: 6.059%; 04: 1.721%; 05: 13.53%; 06: 6.826%; 07: 1.223%; 08: 0.9179%; 09: 1.32%; 10: 2.813%; 11: 2.478%
| PDM3 |  | 01: 2.08%; 03: 0.2454%; 04: -2.315%; 05: -18.75%; 06: -0.9922%; 07: -3.109%; 08: 9.23%; 09: 0.2454%; 10: 0.8002%; 11: 4.086%
| Z-Score (Absolute) |  | 01: 0.64; 03: 0.07; 04: 0.71; 05: 4.51; 06: 0.31; 07: 0.95; 08: 2.83; 09: 0.07; 10: 0.07; 11: 1.25
| Analytical results for Ni in OREAS 905 (Certified Value 9.54 ppm) | [31 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 8.6 | 10.206 | 9.5 | 10 | 9.5 | 8.8 | 8 | 10 | 9.9 | 9 | 8.3 | 9 | 7 | 9.7 | 10 | 9 | 11 | 13.4 | 9.386 | 9 | 11.3 | 10 | 5 | 9 | 9.8 | 11 | 9.3 | 10.3 | 9.8 |  | 1
| 1 | 2 | 9.3 | 9.2279 | 9.6 | 9 | 8.9 | 8.4 | 9 | 10 | 9.7 | 10 | 9.3 | 10 | 7 | 11 | 10 | 8 | 10 | 12.3 | 9.21 | 9 | 10.9 | 9 | 5 | 8.3 | 8.6 | 9 | 8.7 | 8.7 | 9.9 |  | 9
| 1 | 3 | 8.8 | 9.2699 | 9.6 | 10 | 8.6 | 8.7 | 8 | 10 | 10.2 | 9 | 8.9 | 10 | 7 | 9.5 | 11 | 9 | 10 | 12.3 | 9.528 | 9 | 12.3 | 10 | 5 | 11.5 | 8.6 | 11 | 13.3 | 10.1 | 12 |  | 16
| 1 | 4 | 8.6 | 9.3657 | 9.6 | 10 | 9.1 | 8.9 | 10 | 10 | 9.6 | 9 | 8.3 | 10 | 6 | 9.3 | 11 | 9 | 11 | 13.4 | 10.512 | 11 | 11.1 | 10 | 5 | 9.8 | 8.6 | 10 | 8.3 | 9.2 | 13.1 |  | 9.5446
| 1 | 5 | 8.8 | 9.3988 | 9.7 | 10 | 8.7 | 8.5 | 9 | 10 | 9.4 | 9 | 8.8 | 9 | 7 | 9.3 | 11 | 8 | 10 | 13.2 | 9.439 | 9 | 10.7 | 9 | 5 | 9.5 | 8.8 | 10 | 10.2 | 9.3 | 10.1 |  | 41
| 1 | 6 | 9.8 | 9.9846 | 9.7 | 10 | 8.5 | 8.7 | 11 | 10 | 9.9 | 9 | 9.3 | 9 | 8 | 10.3 | 24 | 7 | 10 | 12.9 | 9.36 | 9 | 11.3 | 9 | 5 | 12.9 | 9 | 11 | 9.3 | 9.3 | 11.3
| Mean |  | 8.9833 | 9.5755 | 9.6167 | 9.8333 | 8.8833 | 8.6667 | 9.1667 | 10 | 9.7833 | 9.1667 | 8.8167 | 9.5 | 7 | 9.85 | 12.8333 | 8.3333 | 10.3333 | 12.9167 | 9.5725 | 9.3333 | 11.2667 | 9.5 | 5 | 10.1667 | 8.9 | 10.3333 | 9.85 | 9.4833 | 11.0333
| Median |  | 8.8 | 9.3823 | 9.6 | 10 | 8.8 | 8.7 | 9 | 10 | 9.8 | 9 | 8.85 | 9.5 | 7 | 9.6 | 11 | 8.5 | 10 | 13.05 | 9.4125 | 9 | 11.2 | 9.5 | 5 | 9.65 | 8.7 | 10.5 | 9.3 | 9.3 | 10.7
| Std Dev. |  | 0.47504 | 0.41336 | 0.075277 | 0.40825 | 0.37103 | 0.18619 | 1.169 | 0 | 0.27869 | 0.40825 | 0.44907 | 0.54772 | 0.63246 | 0.67454 | 5.4924 | 0.8165 | 0.5164 | 0.51153 | 0.47198 | 0.8165 | 0.55737 | 0.54772 | 0 | 1.7131 | 0.46904 | 0.8165 | 1.8086 | 0.60139 | 1.3382
| Rel.Std.Dev. |  | 5.288% | 4.317% | 0.7828% | 4.152% | 4.177% | 2.148% | 12.75% | 0% | 2.849% | 4.454% | 5.093% | 5.766% | 9.035% | 6.848% | 42.8% | 9.798% | 4.997% | 3.96% | 4.931% | 8.748% | 4.947% | 5.766% | 0% | 16.85% | 5.27% | 7.902% | 18.36% | 6.342% | 12.13%
| PDM3 |  | -5.881% | 0.3235% | 0.7549% | 3.025% | -6.928% | -9.198% | -3.96% | 4.771% | 2.501% | -3.96% | -7.627% | -0.4675% | -26.66% | 3.2% | 34.46% | -12.69% | 8.263% | 35.33% | 0.2921% | -2.214% | 18.04% | -0.4675% | -47.61% | 6.517% | -6.754% | 8.263% | 3.2% | -0.6421% | 15.6%
| Z-Score (Absolute) |  | 0.92 | 0.1 | 0.16 | 0.45 | 0.83 | 1.12 | 0.45 | 0.67 | 0.38 | 0.45 | 0.92 | 0 | 3.37 | 0.47 | 1.48 | 1.57 | 1.12 | 4.61 | 0.16 | 0.22 | 2.1 | 0 | 6.07 | 0.16 | 1.05 | 1.12 | 0.46 | 0.02 | 1.51
| Analytical results for P in OREAS 905 (Certified Value 0.028 wt.%) | [31 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.025 | 0.026958 | 0.0256 | 0.029 | 0.029 | 0.026 | 0.029 | 0.03 | 0.028 | 0.0276 | 0.024 | 0.026 | 0.03 | 0.025 | < 0.01 | 0.0273 | 0.03 | 0.0272 | 0.025 | 0.031 | 0.022 | 0.025 | 0.027 | 0.026 | 0.02 | 0.026 | 0.0322 | 0.028 |  |  | 1
| 1 | 2 | 0.026 | 0.03057 | 0.0269 | 0.0295 | 0.028 | 0.026 | 0.03 | 0.03 | 0.026 | 0.0285 | 0.025 | 0.0279 | 0.04 | 0.03 | < 0.01 | 0.0271 | 0.03 | 0.0268 | 0.026 | 0.031 | 0.022 | 0.025 | 0.027 | 0.026 | 0.018 | 0.026 | 0.0314 | 0.028 |  |  | 0
| 1 | 3 | 0.028 | 0.029192 | 0.0269 | 0.029 | 0.027 | 0.027 | 0.029 | 0.03 | 0.024 | 0.0266 | 0.026 | 0.0283 | 0.03 | 0.025 | < 0.01 | 0.0267 | 0.03 | 0.0266 | 0.024 | 0.03 | 0.024 | 0.03 | 0.026 | 0.027 | 0.018 | 0.026 | 0.0306 | 0.029 |  |  | 16
| 1 | 4 | 0.027 | 0.027776 | 0.027 | 0.029 | 0.028 | 0.026 | 0.029 | 0.03 | 0.025 | 0.0263 | 0.027 | 0.0271 | 0.03 | 0.028 | 0.02 | 0.027 | 0.03 | 0.027 | 0.023 | 0.032 | 0.022 | 0.03 | 0.027 | 0.026 | 0.02 | 0.027 | 0.0319 | 0.027 |  |  | 0.027505
| 1 | 5 | 0.026 | 0.030214 | 0.0265 | 0.0295 | 0.029 | 0.026 | 0.03 | 0.03 | 0.024 | 0.0276 | 0.025 | 0.0257 | 0.03 | 0.029 | < 0.01 | 0.0274 | 0.03 | 0.0266 | 0.023 | 0.03 | 0.022 | 0.03 | 0.027 | 0.026 | 0.017 | 0.026 | 0.0308 | 0.028 |  |  | 42
| 1 | 6 | 0.026 | 0.02881 | 0.0259 | 0.0315 | 0.029 | 0.026 | 0.029 | 0.03 | 0.025 | 0.0277 | 0.026 | 0.0292 | 0.04 | 0.03 | 0.03 | 0.0272 | 0.03 | 0.0261 | 0.024 | 0.031 | 0.022 | 0.025 | 0.027 | 0.026 | 0.018 | 0.027 | 0.0308 | 0.026
| Mean |  | 0.026333 | 0.02892 | 0.026467 | 0.029583 | 0.028333 | 0.026167 | 0.029333 | 0.03 | 0.025333 | 0.027383 | 0.0255 | 0.027367 | 0.033333 | 0.027833 | 0.025 | 0.027117 | 0.03 | 0.026717 | 0.024167 | 0.030833 | 0.022333 | 0.0275 | 0.026833 | 0.026167 | 0.0185 | 0.026333 | 0.031283 | 0.027667
| Median |  | 0.026 | 0.029001 | 0.0267 | 0.02925 | 0.0285 | 0.026 | 0.029 | 0.03 | 0.025 | 0.0276 | 0.0255 | 0.0275 | 0.03 | 0.0285 | 0.025 | 0.02715 | 0.03 | 0.0267 | 0.024 | 0.031 | 0.022 | 0.0275 | 0.027 | 0.026 | 0.018 | 0.026 | 0.0311 | 0.028
| Std Dev. |  | 0.0010328 | 0.0013885 | 0.00058878 | 0.0009704 | 0.0008165 | 0.00040825 | 0.0005164 | 0 | 0.0015055 | 0.00080353 | 0.0010488 | 0.0013589 | 0.005164 | 0.0023166 | 0.0070711 | 0.00024833 | 0 | 0.00038166 | 0.001169 | 0.00075277 | 0.0008165 | 0.0027386 | 0.00040825 | 0.00040825 | 0.0012247 | 0.0005164 | 0.00065853 | 0.0010328
| Rel.Std.Dev. |  | 3.922% | 4.801% | 2.225% | 3.28% | 2.882% | 1.56% | 1.76% | 0% | 5.943% | 2.934% | 4.113% | 4.966% | 15.49% | 8.323% | 28.28% | 0.9158% | 0% | 1.429% | 4.837% | 2.441% | 3.656% | 9.959% | 1.521% | 1.56% | 6.62% | 1.961% | 2.105% | 3.733%
| PDM3 |  | -4.259% | 5.146% | -3.774% | 7.558% | 3.013% | -4.864% | 6.649% | 9.073% | -7.894% | -0.441% | -7.288% | -0.5016% | 21.19% | 1.195% | -9.106% | -1.411% | 9.073% | -2.865% | -12.14% | 12.1% | -18.8% | -0.0168% | -2.441% | -4.864% | -32.74% | -4.259% | 13.74% | 0.5892%
| Z-Score (Absolute) |  | 0.51 | 0.76 | 0.44 | 0.89 | 0.47 | 0.59 | 0.96 | 1.29 | 1 | 0 | 0.92 | 0 | 2.92 | 0.22 | 1.29 | 0.13 | 1.29 | 0.32 | 1.57 | 1.69 | 2.47 | 0.06 | 0.27 | 0.59 | 4.35 | 0.51 | 1.91 | 0.14
| Analytical results for Pb in OREAS 905 (Certified Value 30.4 ppm) | [31 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 28.1 | 31.2229 | 31.3 | 32 | 30.8 | 28 | 31 | 30 | 29.7 | 32 | 29 | 34 | 27 | 30.1 | 27 | 32 | 23 | 28.3 | 25.9818 | 39 | 34 | 37 | 31 | 34.3 | 28.3 | 36 | 29.9 | 29.5 | 33.16 |  | 1
| 1 | 2 | 27.6 | 30.6555 | 32.2 | 32 | 30 | 28 | 30 | 30 | 29.2 | 33 | 28.2 | 29 | 31 | 30.1 | 25 | 31.9 | 25 | 30.9 | 23.2587 | 35 | 36.6 | 36 | 30 | 29.1 | 28.4 | 32 | 28.5 | 29.7 | 34.39 |  | 0
| 1 | 3 | 28.1 | 31.3103 | 32.3 | 31 | 28.9 | 28.6 | 31 | 30 | 28.8 | 34 | 29 | 36 | 25 | 29.5 | 27 | 31 | 27 | 30.4 | 25.1895 | 38 | 34.4 | 37 | 30 | 34.3 | 28.7 | 31 | 30.1 | 29.6 | 34.68 |  | 16
| 1 | 4 | 27.5 | 31.6055 | 31.6 | 32 | 29.7 | 28.5 | 31 | 30 | 27.1 | 32 | 29 | 31 | 32 | 29.9 | 26 | 31.6 | 21 | 29.5 | 23.575 | 36 | 34.1 | 35 | 29 | 33.4 | 28.8 | 33 | 29.5 | 30.3 | 35.04 |  | 30.4419
| 1 | 5 | 28.2 | 31.7293 | 31.3 | 31 | 30 | 28.4 | 31 | 30 | 28.6 | 30 | 27.3 | 34 | 30 | 30.7 | 26 | 32 | 22 | 28.6 | 23.5435 | 33 | 34.6 | 36 | 31 | 32.1 | 28.8 | 32 | 29.5 | 29.6 | 34.88 |  | 43
| 1 | 6 | 28.1 | 30.8981 | 30.2 | 31 | 29.5 | 28.8 | 31 | 30 | 28.5 | 30 | 28.2 | 30 | 31 | 30.2 | 29 | 31.2 | 25 | 29.9 | 22.5682 | 37 | 33.3 | 43 | 30 | 32.1 | 28.5 | 31 | 30.3 | 29.1 | 35.24
| Mean |  | 27.9333 | 31.2369 | 31.4833 | 31.5 | 29.8167 | 28.3833 | 30.8333 | 30 | 28.65 | 31.8333 | 28.45 | 32.3333 | 29.3333 | 30.0833 | 26.6667 | 31.6167 | 23.8333 | 29.6 | 24.0194 | 36.3333 | 34.5 | 37.3333 | 30.1667 | 32.55 | 28.5833 | 32.5 | 29.6333 | 29.6333 | 34.565
| Median |  | 28.1 | 31.2666 | 31.45 | 31.5 | 29.85 | 28.45 | 31 | 30 | 28.7 | 32 | 28.6 | 32.5 | 30.5 | 30.1 | 26.5 | 31.75 | 24 | 29.7 | 23.5592 | 36.5 | 34.25 | 36.5 | 30 | 32.75 | 28.6 | 32 | 29.7 | 29.6 | 34.78
| Std Dev. |  | 0.30111 | 0.40909 | 0.76267 | 0.54772 | 0.63061 | 0.32506 | 0.40825 | 0 | 0.87807 | 1.6021 | 0.68629 | 2.7325 | 2.7325 | 0.392 | 1.3663 | 0.43089 | 2.2286 | 1.0119 | 1.2906 | 2.1602 | 1.1207 | 2.8752 | 0.75277 | 1.9573 | 0.2137 | 1.8708 | 0.64083 | 0.38816 | 0.74811
| Rel.Std.Dev. |  | 1.078% | 1.31% | 2.422% | 1.739% | 2.115% | 1.145% | 1.324% | 0% | 3.065% | 5.033% | 2.412% | 8.451% | 9.315% | 1.303% | 5.123% | 1.363% | 9.351% | 3.419% | 5.373% | 5.946% | 3.248% | 7.701% | 2.495% | 6.013% | 0.7476% | 5.756% | 2.163% | 1.31% | 2.164%
| PDM3 |  | -8.24% | 2.612% | 3.421% | 3.476% | -2.054% | -6.762% | 1.286% | -1.451% | -5.886% | 4.571% | -6.543% | 6.213% | -3.641% | -1.178% | -12.4% | 3.859% | -21.71% | -2.765% | -21.1% | 19.35% | 13.33% | 22.64% | -0.904% | 6.925% | -6.105% | 6.761% | -2.656% | -2.656% | 13.54%
| Z-Score (Absolute) |  | 0.95 | 0.51 | 0.62 | 0.62 | 0.12 | 0.75 | 0.33 | 0.04 | 0.63 | 0.77 | 0.72 | 0.99 | 0.33 | 0 | 1.5 | 0.67 | 2.75 | 0.21 | 2.67 | 2.75 | 1.76 | 2.69 | 0.04 | 1.08 | 0.66 | 0.75 | 0.2 | 0.2 | 2.09
| Analytical results for Pr in OREAS 905 (Certified Value 10.5 ppm) | [31 empty] | Certified
Pr | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 17; 08: 20; 09: 26; 10: 30; 11: 33 |  | ppm
|  |  | 01: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*OES/MS |  | 1
2
| 1 | 1 | 01: 9.54; 03: 10.4; 04: 10.7; 05: 9.9; 06: 9.3; 07: 10.56; 08: 11.205; 09: 10.8; 10: 11.1; 11: 11.3 |  | 1
| 1 | 2 | 01: 10; 03: 10.1; 04: 10.6; 05: 9.2; 06: 9.1; 07: 10.59; 08: 11.241; 09: 10.7; 10: 11.1; 11: 11 |  | 0
| 1 | 3 | 01: 10; 03: 10.1; 04: 10.5; 05: 8.9; 06: 10.5; 07: 10.32; 08: 11.473; 09: 10.7; 10: 11.2; 11: 10.9 |  | 16
| 1 | 4 | 01: 10.3; 03: 10; 04: 10.2; 05: 9.8; 06: 10.4; 07: 10.5; 08: 11.484; 09: 10.8; 10: 12; 11: 11.1 |  | 10.5396
| 1 | 5 | 01: 10.1; 03: 10.6; 04: 10.6; 05: 6.3; 06: 9.9; 07: 10.73; 08: 11.541; 09: 10.8; 10: 11.2; 11: 11.5 |  | 44
| 1 | 6 | 01: 10; 03: 10.5; 04: 10.6; 05: 8.4; 06: 11.1; 07: 10.63; 08: 11.379; 09: 10.7; 10: 11.3; 11: 11.5
| Mean |  | 01: 9.99; 03: 10.2833; 04: 10.5333; 05: 8.75; 06: 10.05; 07: 10.555; 08: 11.3872; 09: 10.75; 10: 11.3167; 11: 11.2167
| Median |  | 01: 10; 03: 10.25; 04: 10.6; 05: 9.05; 06: 10.15; 07: 10.575; 08: 11.426; 09: 10.75; 10: 11.2; 11: 11.2
| Std Dev. |  | 01: 0.2494; 03: 0.24833; 04: 0.17512; 05: 1.3248; 06: 0.76354; 07: 0.13838; 08: 0.13786; 09: 0.054772; 10: 0.34303; 11: 0.25626
| Rel.Std.Dev. |  | 01: 2.496%; 03: 2.415%; 04: 1.663%; 05: 15.14%; 06: 7.597%; 07: 1.311%; 08: 1.211%; 09: 0.5095%; 10: 3.031%; 11: 2.285%
| PDM3 |  | 01: -5.214%; 03: -2.431%; 04: -0.05898%; 05: -16.98%; 06: -4.645%; 07: 0.1466%; 08: 8.042%; 09: 1.997%; 10: 7.373%; 11: 6.425%
| Z-Score (Absolute) |  | 01: 0.71; 03: 0.34; 04: 0.01; 05: 1.41; 06: 0.64; 07: 0.01; 08: 1.08; 09: 0.26; 10: 0.82; 11: 0.87
| Analytical results for Rb in OREAS 905 (Certified Value 138 ppm) | [31 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 17 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 0
0
| 1 | 1 | 131.5 | 144.3523 | 133.7 | 138 | 136.5 | 141 | 130 | 145 | 151 | 127 | 108.2 | 137 | 140.73 | 141.7688 | 141.8 | 147 | 133.5 | 129.5 | 163 | 137 | 144 | 133.7 | [8 empty] | 1
| 1 | 2 | 131.5 | 145.7599 | 133.2 | 144 | 133 | 141 | 133 | 148 | 151 | 122 | 114.3 | 139 | 136.4 | 142.8216 | 145.2 | 140 | 124 | 132.5 | 165 | 137 | 144 | 137.4 | [8 empty] | 0
| 1 | 3 | 138 | 146.6644 | 134.8 | 142 | 130.5 | 144 | 130 | 146 | 140 | 117 | 136.4 | 137 | 137.63 | 142.167 | 141.8 | 118 | 135 | 135 | 174 | 139 | 145 | 137.3 | [8 empty] | 16
| 1 | 4 | 129.5 | 145.6432 | 132.1 | 145 | 134.5 | 142.5 | 133 | 140 | 136 | 126 | 137.8 | 139 | 137.22 | 138.5288 | 141.3 | 148 | 134.5 | 135 | 172 | 140 | 145 | 136.3 | [8 empty] | 138.2814
| 1 | 5 | 130.5 | 143.8627 | 134.5 | 144 | 135.5 | 145.5 | 134 | 151 | 142 | 124 | 130.3 | 141 | 140.61 | 140.4943 | 146.4 | 143 | 138 | 133 | 174 | 139 | 143 | 135.4 | [8 empty] | 45
| 1 | 6 | 133.5 | 144.8312 | 130.6 | 145 | 133 | 142.5 | 131 | 145 | 145 | 121 | 136.7 | 140 | 140.22 | 139.8488 | 143.3 | 146 | 138 | 132 | 169 | 145.5 | 140 | 139
| Mean |  | 132.4167 | 145.1856 | 133.15 | 143 | 133.8333 | 142.75 | 131.8333 | 145.8333 | 144.1667 | 122.8333 | 127.2833 | 138.8333 | 138.8017 | 140.9382 | 143.3 | 140.3333 | 133.8333 | 132.8333 | 169.5 | 139.5833 | 143.5 | 136.5167
| Median |  | 131.5 | 145.2372 | 133.45 | 144 | 133.75 | 142.5 | 132 | 145.5 | 143.5 | 123 | 133.35 | 139 | 138.925 | 141.1316 | 142.55 | 144.5 | 134.75 | 132.75 | 170.5 | 139 | 144 | 136.8
| Std Dev. |  | 3.04 | 1.0292 | 1.5783 | 2.6833 | 2.137 | 1.7536 | 1.7224 | 3.656 | 6.047 | 3.656 | 12.8392 | 1.6021 | 1.9309 | 1.6085 | 2.0842 | 11.3255 | 5.164 | 2.0656 | 4.6797 | 3.1371 | 1.8708 | 1.8324
| Rel.Std.Dev. |  | 2.296% | 0.7089% | 1.185% | 1.876% | 1.597% | 1.228% | 1.306% | 2.507% | 4.194% | 2.976% | 10.09% | 1.154% | 1.391% | 1.141% | 1.454% | 8.07% | 3.859% | 1.555% | 2.761% | 2.248% | 1.304% | 1.342%
| PDM3 |  | -4.241% | 4.993% | -3.711% | 3.412% | -3.217% | 3.232% | -4.663% | 5.461% | 4.256% | -11.17% | -7.953% | 0.3992% | 0.3763% | 1.921% | 3.629% | 1.484% | -3.217% | -3.94% | 22.58% | 0.9415% | 3.774% | -1.276%
| Z-Score (Absolute) |  | 1 | 0.84 | 0.75 | 0.69 | 0.66 | 0.52 | 0.93 | 0.93 | 0.71 | 2.12 | 0.47 | 0 | 0 | 0.28 | 0.59 | 0.79 | 0.4 | 0.79 | 4.07 | 0.06 | 0.62 | 0.31
| Analytical results for Re in OREAS 905 (Indicative Value < 0.002 ppm) | [31 empty] | Indicative
Re | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab; 17: Lab; 18: Lab |  | 1
| No | No | 01: 01; 03: 03; 04: 06; 05: 07; 06: 08; 07: 09; 08: 11; 09: 12; 10: 13; 11: 17; 12: 20; 13: 24; 14: 26; 15: 28; 16: 29; 17: 31; 18: 33 |  | ppm
|  |  | 01: 4A*OES/MS; 03: 4A*OES; 04: 4A*OES/MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*OES/MS; 14: 4A*MS; 15: 4A*OES/MS; 16: 4A*OES/MS; 17: 4A*OES/MS; 18: 4A*OES/MS |  | 3
3
| 1 | 1 | 01: < 0.002; 03: < 0.05; 04: 0.002; 05: < 0.002; 06: < 0.05; 07: < 0.1; 08: < 5; 09: < 0.001; 10: < 0.05; 11: < 0.002; 12: < 0.05; 13: 0.536; 14: < 0.1; 15: < 0.002; 16: < 0.002; 17: < 0.002; 18: < 0.002 |  | 1
| 1 | 2 | 01: < 0.002; 03: < 0.05; 04: 0.002; 05: < 0.002; 06: < 0.05; 07: < 0.1; 08: < 5; 09: < 0.001; 10: < 0.05; 11: < 0.002; 12: < 0.05; 13: 0.051; 14: < 0.1; 15: < 0.002; 16: < 0.002; 17: < 0.002; 18: < 0.002 |  | 4
| 1 | 3 | 01: < 0.002; 03: < 0.05; 04: 0.002; 05: < 0.002; 06: < 0.05; 07: < 0.1; 08: < 5; 09: < 0.001; 10: < 0.05; 11: < 0.002; 12: < 0.05; 13: 0.034; 14: < 0.1; 15: 0.003; 16: < 0.002; 17: < 0.002; 18: < 0.002 |  | 16
| 1 | 4 | 01: < 0.002; 03: < 0.05; 04: 0.002; 05: < 0.002; 06: < 0.05; 07: < 0.1; 08: < 5; 09: < 0.001; 10: < 0.05; 11: < 0.002; 12: < 0.05; 13: 0.018; 14: < 0.1; 15: < 0.002; 16: < 0.002; 17: < 0.002; 18: < 0.002 |  | < 0.002
| 1 | 5 | 01: < 0.002; 03: < 0.05; 04: 0.002; 05: < 0.002; 06: < 0.05; 07: < 0.1; 08: < 5; 09: < 0.001; 10: < 0.05; 11: < 0.002; 12: < 0.05; 13: 0.01; 14: < 0.1; 15: < 0.002; 16: < 0.002; 17: < 0.002; 18: 0.003 |  | 10
| 1 | 6 | 01: < 0.002; 03: < 0.05; 04: < 0.002; 05: < 0.002; 06: < 0.05; 07: < 0.1; 08: < 5; 09: < 0.001; 10: < 0.05; 11: < 0.002; 12: < 0.05; 13: 0.012; 14: < 0.1; 15: 0.002; 16: < 0.002; 17: < 0.002; 18: 0.003
| Mean |  | 04: 0.002; 13: 0.11017; 15: 0.0025; 18: 0.003
| Median |  | 04: 0.002; 13: 0.026; 15: 0.0025; 18: 0.003
| Std Dev. |  | 04: 0; 13: 0.20919; 15: 0.00070711; 18: 0
| Rel.Std.Dev. |  | 04: 0%; 13: 189.9%; 15: 28.28%; 18: 0%
PDM3
| Z-Score (Absolute) |  | 01: NA; 03: NA; 04: 0.34; 05: NA; 06: NA; 07: NA; 08: NA; 09: NA; 10: NA; 11: NA; 12: NA; 13: 94.07; 14: NA; 15: 1.01; 16: NA; 17: NA; 18: 0.34
| Analytical results for S in OREAS 905 (Certified Value 0.066 wt.%) | [31 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 3
3
| 1 | 1 | 0.06 | 0.054793 | 0.08 | 0.067 | 0.07 | 0.07 | 0.056 | 0.07 | 0.08 | 0.07 | 0.06 | 0.067 | 0.07 | 0.13 | 0.0626 | 0.07 | 0.074 | 0.062982 | 0.1 | 0.06 | < 0.1 | 0.065 | 0.07 | 0.06 | 0.076 | 0.06 | 0.07 | 0.06 |  |  | 1
| 1 | 2 | 0.07 | 0.05414 | 0.08 | 0.0665 | 0.07 | 0.06 | 0.054 | 0.075 | 0.08 | 0.07 | 0.05 | 0.0544 | 0.07 | 0.08 | 0.08 | 0.07 | 0.063 | 0.070394 | 0.1 | 0.06 | < 0.1 | 0.07 | 0.07 | 0.07 | 0.076 | 0.07 | 0.06 | 0.05 |  |  | 24
| 1 | 3 | 0.07 | 0.05488 | 0.08 | 0.065 | 0.06 | 0.07 | 0.053 | 0.07 | 0.08 | 0.07 | 0.06 | 0.0651 | 0.07 | 0.08 | 0.0632 | 0.06 | 0.063 | 0.07101 | < 0.1 | 0.06 | < 0.1 | 0.065 | 0.07 | 0.07 | 0.078 | 0.07 | 0.07 | 0.05 |  |  | 16
| 1 | 4 | 0.07 | 0.054498 | 0.08 | 0.067 | 0.07 | 0.06 | 0.055 | 0.07 | 0.07 | 0.07 | 0.06 | 0.058 | 0.07 | 0.1 | 0.061 | 0.07 | 0.063 | 0.065123 | < 0.1 | 0.06 | < 0.1 | 0.065 | 0.07 | 0.07 | 0.077 | 0.07 | 0.07 | 0.05 |  |  | 0.066039
| 1 | 5 | 0.07 | 0.055258 | 0.09 | 0.0665 | 0.07 | 0.07 | 0.055 | 0.075 | 0.07 | 0.07 | 0.06 | 0.0607 | 0.07 | 0.08 | 0.0627 | 0.07 | 0.074 | 0.075046 | < 0.1 | 0.06 | < 0.1 | 0.065 | 0.06 | 0.07 | 0.074 | 0.06 | 0.06 | 0.05 |  |  | 46
| 1 | 6 | 0.07 | 0.055138 | 0.1 | 0.0685 | 0.07 | 0.07 | 0.054 | 0.07 | 0.07 | 0.07 | 0.06 | 0.0598 | 0.07 | 0.1 | 0.0629 | 0.07 | 0.063 | 0.066774 | < 0.1 | 0.06 | < 0.1 | 0.065 | 0.07 | 0.06 | 0.08 | 0.06 | 0.06 | 0.05
| Mean |  | 0.068333 | 0.054785 | 0.085 | 0.06675 | 0.068333 | 0.066667 | 0.0545 | 0.071667 | 0.075 | 0.07 | 0.058333 | 0.060833 | 0.07 | 0.095 | 0.0654 | 0.068333 | 0.066667 | 0.068555 | 0.1 | 0.06 |  | 0.065833 | 0.068333 | 0.066667 | 0.076833 | 0.065 | 0.065 | 0.051667
| Median |  | 0.07 | 0.054837 | 0.08 | 0.06675 | 0.07 | 0.07 | 0.0545 | 0.07 | 0.075 | 0.07 | 0.06 | 0.06025 | 0.07 | 0.09 | 0.0628 | 0.07 | 0.063 | 0.068584 | 0.1 | 0.06 |  | 0.065 | 0.07 | 0.07 | 0.0765 | 0.065 | 0.065 | 0.05
| Std Dev. |  | 0.0040825 | 0.00041365 | 0.0083666 | 0.0011292 | 0.0040825 | 0.005164 | 0.0010488 | 0.002582 | 0.0054772 | 0 | 0.0040825 | 0.0046202 | 0 | 0.019748 | 0.0071936 | 0.0040825 | 0.0056804 | 0.0044169 | 0 | 0 |  | 0.0020412 | 0.0040825 | 0.005164 | 0.0020412 | 0.0054772 | 0.0054772 | 0.0040825
| Rel.Std.Dev. |  | 5.974% | 0.755% | 9.843% | 1.692% | 5.974% | 7.746% | 1.924% | 3.603% | 7.303% | 0% | 6.999% | 7.595% | 0% | 20.79% | 11% | 5.974% | 8.521% | 6.443% | 0% | 0% |  | 3.101% | 5.974% | 7.746% | 2.657% | 8.427% | 8.427% | 7.902%
| PDM3 |  | 3.474% | -17.04% | 28.71% | 1.076% | 3.474% | 0.9501% | -17.47% | 8.521% | 13.57% | 5.998% | -11.67% | -7.883% | 5.998% | 43.85% | -0.968% | 3.474% | 0.9501% | 3.809% | 51.43% | -9.145% |  | -0.3118% | 3.474% | 0.9501% | 16.34% | -1.574% | -1.574% | -21.76%
| Z-Score (Absolute) |  | 0.33 | 2.48 | 3.79 | 0 | 0.33 | 0.02 | 2.54 | 1.02 | 1.71 | 0.67 | 1.75 | 1.23 | 0.67 | 5.86 | 0.89 | 0.33 | 0.02 | 0.37 | 6.9 | 1.4 | NA | 0.19 | 0.33 | 0.02 | 2.09 | 0.36 | 0.36 | 3.13
| Analytical results for Sb in OREAS 905 (Certified Value 1.95 ppm) | [31 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 1.91 | 1.8881 | 2 | 1.9 | 1.95 | 1.98 | 1.9 | 1.7 | 1.15 | < 5 | 1.9 | 1.9 | < 10 | 2.15 | < 2 | 2.29 | < 2 | < 2 | 2.305 | < 5 | 6.4 | < 5 | 2 | 1.85 | 1.69 | 3 | 1.92 | 1.99 | 2 |  | 1
| 1 | 2 | 1.91 | 1.9818 | 2.01 | 1.9 | 1.91 | 2.03 | 1.9 | 1.8 | 1.02 | < 5 | 2.1 | 1.8 | < 10 | 2.02 | < 2 | 2.18 | < 2 | < 2 | 2.226 | < 5 | 4.2 | < 5 | 2 | 1.69 | 1.76 | 2 | 1.88 | 1.96 | 2.07 |  | 0
| 1 | 3 | 1.99 | 2.0394 | 1.93 | 1.9 | 1.85 | 2.06 | 1.9 | 1.7 | 0.88 | < 5 | 1.7 | 1.9 | < 10 | 2.02 | < 2 | 2.26 | < 2 | < 2 | 2.092 | < 5 | 3.6 | < 5 | 2 | 2.21 | 1.79 | 3 | 1.91 | 2.06 | 2.13 |  | 16
| 1 | 4 | 1.9 | 1.916 | 1.91 | 2 | 1.92 | 2.03 | 2 | 1.6 | 0.82 | < 5 | 2 | 1.8 | < 10 | 2 | < 2 | 2.17 | < 2 | < 2 | 2.258 | < 5 | 3.1 | < 5 | 2 | 2.18 | 1.74 | < 2 | 1.92 | 2.1 | 2.11 |  | 1.9516
| 1 | 5 | 1.94 | 1.9126 | 1.96 | 2 | 1.93 | 2.07 | 2 | 1.8 | 1.14 | < 5 | 2 | 1.7 | < 10 | 2.04 | < 2 | 2.29 | < 2 | < 2 | 1.944 | < 5 | 2.9 | < 5 | 2 | 1.86 | 1.76 | 3 | 1.94 | 2.12 | 2.09 |  | 47
| 1 | 6 | 1.97 | 2.018 | 1.86 | 1.9 | 1.86 | 2.09 | 2 | 1.8 | 0.91 | < 5 | 2 | 1.8 | < 10 | 2.07 | < 2 | 2.34 | < 2 | < 2 | 1.998 | < 5 | 2.7 | < 5 | 2 | 1.89 | 1.7 | 4 | 1.96 | 1.95 | 2.21
| Mean |  | 1.9367 | 1.9593 | 1.945 | 1.9333 | 1.9033 | 2.0433 | 1.95 | 1.7333 | 0.98667 |  | 1.95 | 1.8167 |  | 2.05 |  | 2.255 |  |  | 2.1372 |  | 3.8167 |  | 2 | 1.9467 | 1.74 | 3 | 1.9217 | 2.03 | 2.1017
| Median |  | 1.925 | 1.9489 | 1.945 | 1.9 | 1.915 | 2.045 | 1.95 | 1.75 | 0.965 |  | 2 | 1.8 |  | 2.03 |  | 2.275 |  |  | 2.159 |  | 3.35 |  | 2 | 1.875 | 1.75 | 3 | 1.92 | 2.025 | 2.1
| Std Dev. |  | 0.036697 | 0.062434 | 0.056833 | 0.05164 | 0.039833 | 0.038816 | 0.054772 | 0.08165 | 0.1388 |  | 0.13784 | 0.075277 |  | 0.054406 |  | 0.067157 |  |  | 0.14791 |  | 1.3761 |  | 0 | 0.20481 | 0.038471 | 0.70711 | 0.027142 | 0.073212 | 0.069402
| Rel.Std.Dev. |  | 1.895% | 3.187% | 2.922% | 2.671% | 2.093% | 1.9% | 2.809% | 4.711% | 14.07% |  | 7.069% | 4.144% |  | 2.654% |  | 2.978% |  |  | 6.921% |  | 36.06% |  | 0% | 10.52% | 2.211% | 23.57% | 1.412% | 3.607% | 3.302%
| PDM3 |  | -0.7633% | 0.3971% | -0.3363% | -0.9341% | -2.471% | 4.702% | -0.08013% | -11.18% | -49.44% |  | -0.08013% | -6.912% |  | 5.044% |  | 15.55% |  |  | 9.51% |  | 95.57% |  | 2.482% | -0.2509% | -10.84% | 53.72% | -1.532% | 4.019% | 7.691%
| Z-Score (Absolute) |  | 0.16 | 0.04 | 0.09 | 0.19 | 0.46 | 0.79 | 0.04 | 1.98 | 8.66 | NA | 0.41 | 1.23 | NA | 0.67 | NA | 2.69 | NA | NA | 1.63 | NA | 12.04 | NA | 0.41 | 0.07 | 1.92 | 9.36 | 0.3 | 0.67 | 1.32
| Analytical results for Sc in OREAS 905 (Certified Value 4.9 ppm) | [31 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 2
3
| 1 | 1 | 5.5 | 4.19 | 4.9 | 4.5 | 5 | 5.2 | 5 | 5 | 5.3 | 5 | 6 | 4 | < 5 | 5.3 | 4 | 5 | 3 | 4.8 | 4.7928 | 5 | 7.1 | 4 | 5 | 5.3 | 5.2 | 5.1 | 5 | 5.1 | 4.5 |  | 1
| 1 | 2 | 5.8 | 4.3064 | 5.1 | 4.7 | 4.8 | 5.2 | 5 | 5 | 5.2 | 5 | 5 | 4 | < 5 | 5 | 4 | 5 | 3 | 4.6 | 4.7736 | 5 | 7.4 | 4 | 5 | 4.7 | 5.4 | 5 | 4.9 | 5.1 | 4.9 |  | 6
| 1 | 3 | 5.9 | 3.8611 | 5.3 | 4.8 | 4.7 | 5.3 | 5 | 5 | 5.2 | 5 | 6 | 5 | < 5 | 5 | 4 | 5 | 2 | 4.5 | 4.7096 | 4 | 6.5 | 4 | 5 | 5.1 | 5.4 | 5 | 5 | 5.2 | 4.9 |  | 16
| 1 | 4 | 5.7 | 4.0006 | 5 | 4.9 | 4.9 | 5.2 | 5 | 5 | 5 | 5 | 6 | 5 | < 5 | 5 | 4 | 5 | 3 | 4.4 | 4.7489 | 4 | 6.6 | 4 | 5 | 5.2 | 5.4 | 5.3 | 4.9 | 5.3 | 4.6 |  | 4.8952
| 1 | 5 | 5.7 | 4.086 | 4.9 | 5.2 | 4.8 | 5.2 | 5 | 5 | 5.2 | 5 | 5 | 4 | < 5 | 5 | 4 | 5 | 3 | 5.1 | 4.6926 | 4 | 6.9 | 5 | 5 | 5.2 | 5.5 | 5.1 | 5 | 5.2 | 4.8 |  | 48
| 1 | 6 | 5.6 | 4.1226 | 5 | 4.6 | 4.7 | 5.2 | 5 | 5 | 5.2 | 5 | 5 | 5 | < 5 | 5 | 4 | 5 | 3 | 4.6 | 4.6519 | 4 | 6.5 | 4 | 5 | 5.3 | 5.2 | 4.8 | 5.2 | 5.2 | 4.8
| Mean |  | 5.7 | 4.0945 | 5.0333 | 4.7833 | 4.8167 | 5.2167 | 5 | 5 | 5.1833 | 5 | 5.5 | 4.5 |  | 5.05 | 4 | 5 | 2.8333 | 4.6667 | 4.7282 | 4.3333 | 6.8333 | 4.1667 | 5 | 5.1333 | 5.35 | 5.05 | 5 | 5.1833 | 4.75
| Median |  | 5.7 | 4.1043 | 5 | 4.75 | 4.8 | 5.2 | 5 | 5 | 5.2 | 5 | 5.5 | 4.5 |  | 5 | 4 | 5 | 3 | 4.6 | 4.7292 | 4 | 6.75 | 4 | 5 | 5.2 | 5.4 | 5.05 | 5 | 5.2 | 4.8
| Std Dev. |  | 0.14142 | 0.15375 | 0.15055 | 0.24833 | 0.1169 | 0.040825 | 0 | 0 | 0.098319 | 0 | 0.54772 | 0.54772 |  | 0.12247 | 0 | 0 | 0.40825 | 0.25033 | 0.053093 | 0.5164 | 0.36697 | 0.40825 | 0 | 0.22509 | 0.12247 | 0.16432 | 0.10954 | 0.075277 | 0.16432
| Rel.Std.Dev. |  | 2.481% | 3.755% | 2.991% | 5.192% | 2.427% | 0.7826% | 0% | 0% | 1.897% | 0% | 9.959% | 12.17% |  | 2.425% | 0% | 0% | 14.41% | 5.364% | 1.123% | 11.92% | 5.37% | 9.798% | 0% | 4.385% | 2.289% | 3.254% | 2.191% | 1.452% | 3.459%
| PDM3 |  | 16.44% | -16.36% | 2.821% | -2.286% | -1.605% | 6.566% | 2.14% | 2.14% | 5.885% | 2.14% | 12.35% | -8.074% |  | 3.162% | -18.29% | 2.14% | -42.12% | -4.669% | -3.412% | -11.48% | 39.59% | -14.88% | 2.14% | 4.864% | 9.29% | 3.162% | 2.14% | 5.885% | -2.967%
| Z-Score (Absolute) |  | 2.16 | 2.8 | 0.1 | 0.67 | 0.57 | 0.67 | 0 | 0 | 0.57 | 0 | 1.54 | 1.54 | NA | 0 | 3.09 | 0 | 6.69 | 1.03 | 0.84 | 2.06 | 5.66 | 2.57 | 0 | 0.68 | 1.08 | 0.15 | 0 | 0.57 | 0.77
| Analytical results for Se in OREAS 905 (Certified Value 2.84 ppm) | [31 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 20 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 3 | 2.6272 | < 2 | 3 | 3 | 3 | 3 | < 5 | < 2 | < 5 | 5.2 | 3 | < 20 | 2 | < 12 | 2.6 | < 12 | 2.4761 | < 5 | 3 | 3 | < 2 | 3 | 3 | 1.8 | [5 empty] | 1
| 1 | 2 | 3 | 2.263 | < 2 | 3 | 3 | 3 | 3 | < 5 | < 2 | < 5 | 4.8 | 2 | < 20 | 3 | < 12 | 2.5 | < 12 | 2.8671 | < 5 | 3 | 3 | < 2 | 3 | 3 | 2.1 | [5 empty] | 7
| 1 | 3 | 3 | 2.1188 | < 2 | 3 | 3 | 3 | 3 | < 5 | < 2 | < 5 | 5.7 | 3 | < 20 | 3 | < 12 | 2.5 | < 12 | 2.2415 | < 5 | 3 | 3 | < 2 | 3 | 3 | 2 | [5 empty] | 16
| 1 | 4 | 3 | 2.4669 | < 2 | 2 | 3 | 3 | 3 | < 5 | < 2 | < 5 | 5.8 | 2 | < 20 | 3 | < 12 | 2.8 | < 12 | 3.3428 | < 5 | 3 | 4 | < 2 | 3 | 3 | 1.8 | [5 empty] | 2.8408
| 1 | 5 | 3 | 2.5353 | < 2 | 3 | 3 | 3 | 2 | < 5 | < 2 | < 5 | 4 | 3 | < 20 | 3 | < 12 | 3 | < 12 | 3.0354 | < 5 | 3 | 3 | < 2 | 3 | 3 | 2.4 | [5 empty] | 49
| 1 | 6 | 3 | 2.257 | < 2 | 3 | 3 | 3 | 3 | < 5 | < 2 | < 5 | 5 | 2 | < 20 | 3 | < 12 | 2.5 | < 12 | 2.9163 | < 5 | 3 | 3 | < 2 | 3 | 3 | 1.8
| Mean |  | 3 | 2.378 |  | 2.8333 | 3 | 3 | 2.8333 |  |  |  | 5.0833 | 2.5 |  | 2.8333 |  | 2.65 |  | 2.8132 |  | 3 | 3.1667 |  | 3 | 3 | 1.9833
| Median |  | 3 | 2.365 |  | 3 | 3 | 3 | 3 |  |  |  | 5.1 | 2.5 |  | 3 |  | 2.55 |  | 2.8917 |  | 3 | 3 |  | 3 | 3 | 1.9
| Std Dev. |  | 0 | 0.19482 |  | 0.40825 | 0 | 0 | 0.40825 |  |  |  | 0.65853 | 0.54772 |  | 0.40825 |  | 0.20736 |  | 0.39596 |  | 0 | 0.40825 |  | 0 | 0 | 0.24014
| Rel.Std.Dev. |  | 0% | 8.192% |  | 14.41% | 0% | 0% | 14.41% |  |  |  | 12.95% | 21.91% |  | 14.41% |  | 7.825% |  | 14.08% |  | 0% | 12.89% |  | 0% | 0% | 12.11%
| PDM3 |  | 5.604% | -16.29% |  | -0.263% | 5.604% | 5.604% | -0.263% |  |  |  | 78.94% | -12% |  | -0.263% |  | -6.717% |  | -0.9711% |  | 5.604% | 11.47% |  | 5.604% | 5.604% | -30.18%
| Z-Score (Absolute) |  | 0.67 | 4.36 | NA | 0.67 | 0.67 | 0.67 | 0.67 | NA | NA | NA | 17.53 | 3.37 | NA | 0.67 | NA | 2.72 | NA | 0.84 | NA | 0.67 | 0.67 | NA | 0.67 | 0.67 | 7.55
| Analytical results for Sm in OREAS 905 (Certified Value 7.4 ppm) | [31 empty] | Certified
Sm | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab | [5 empty] | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 17; 08: 20; 09: 26; 10: 30; 11: 33 | [5 empty] | ppm
|  |  | 01: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 01: 7.2; 03: 7.45; 04: 7.1; 05: 6.5; 06: 7; 07: 7.82; 08: 7.765; 09: 7.65; 10: 6.9; 11: 8.2 | [5 empty] | 1
| 1 | 2 | 01: 7.7; 03: 7.75; 04: 7.1; 05: 6.3; 06: 6.9; 07: 7.49; 08: 8.077; 09: 7.55; 10: 6.9; 11: 7.8 | [5 empty] | 0
| 1 | 3 | 01: 7.5; 03: 7.55; 04: 7; 05: 6; 06: 7.9; 07: 7.46; 08: 8.043; 09: 7.45; 10: 7; 11: 8 | [5 empty] | 16
| 1 | 4 | 01: 7.7; 03: 7.35; 04: 6.8; 05: 6.5; 06: 7.8; 07: 7.42; 08: 8.308; 09: 7.6; 10: 7.3; 11: 7.7 | [5 empty] | 7.3993
| 1 | 5 | 01: 7.7; 03: 7.9; 04: 7.1; 05: 4.5; 06: 7.7; 07: 7.89; 08: 8.199; 09: 7.7; 10: 6.8; 11: 7.1 | [5 empty] | 50
| 1 | 6 | 01: 7.6; 03: 7.8; 04: 7.1; 05: 5.6; 06: 8.2; 07: 7.7; 08: 7.847; 09: 7.4; 10: 6.9; 11: 7.1
| Mean |  | 01: 7.5667; 03: 7.6333; 04: 7.0333; 05: 5.9; 06: 7.5833; 07: 7.63; 08: 8.0398; 09: 7.5583; 10: 6.9667; 11: 7.65
| Median |  | 01: 7.65; 03: 7.65; 04: 7.1; 05: 6.15; 06: 7.75; 07: 7.595; 08: 8.06; 09: 7.575; 10: 6.9; 11: 7.75
| Std Dev. |  | 01: 0.19664; 03: 0.21602; 04: 0.12111; 05: 0.76681; 06: 0.51929; 07: 0.2006; 08: 0.2056; 09: 0.11583; 10: 0.17512; 11: 0.45935
| Rel.Std.Dev. |  | 01: 2.599%; 03: 2.83%; 04: 1.722%; 05: 13%; 06: 6.848%; 07: 2.629%; 08: 2.557%; 09: 1.532%; 10: 2.514%; 11: 6.005%
| PDM3 |  | 01: 2.262%; 03: 3.163%; 04: -4.946%; 05: -20.26%; 06: 2.487%; 07: 3.118%; 08: 8.656%; 09: 2.149%; 10: -5.847%; 11: 3.388%
| Z-Score (Absolute) |  | 01: 0.49; 03: 0.39; 04: 8.43; 05: 18.86; 06: 0.34; 07: 0.34; 08: 6.37; 09: 0.71; 10: 10.4; 11: 0.64
| Analytical results for Sn in OREAS 905 (Certified Value 3.96 ppm) | [31 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 4 | 4.1826 | 3.5 | 4 | 3.7 | 3.9 | 6.4 | 4.1 | 3.7 | < 10 | 3 | 3.8 | < 20 | 4.3 | < 5 | 4.3 | < 5 | < 5 | 6 | < 0.1 | 3 | 4 | 3.9 | 3.9 | 2 | 3.9 | 3.9 | 4.2 |  |  | 1
| 1 | 2 | 4.1 | 4.1902 | 3.9 | 4.3 | 3.7 | 4 | 6.2 | 4.2 | 3.6 | < 10 | 4 | 4 | < 20 | 3.8 | < 5 | 4.2 | 5 | < 5 | 4 | 5.2 | 4 | 4 | 3.4 | 4 | < 2 | 3.8 | 3.9 | 4.2 |  |  | 0
| 1 | 3 | 4.1 | 4.1964 | 4 | 4 | 3.6 | 4.2 | 6.3 | 4.2 | 3.4 | < 10 | 4 | 4.2 | < 20 | 3.8 | < 5 | 4 | < 5 | < 5 | 4 | < 0.1 | 3 | 4 | 4.2 | 4.1 | 3 | 4 | 4 | 4.1 |  |  | 16
| 1 | 4 | 3.9 | 4.3369 | 3.8 | 4.5 | 3.7 | 5 | 6.4 | 4.3 | 3.4 | < 10 | 4 | 4.1 | < 20 | 4 | < 5 | 4.1 | 5 | < 5 | 2 | < 0.1 | 3 | 4 | 4.2 | 4 | 2 | 3.9 | 3.9 | 4.4 |  |  | 3.9563
| 1 | 5 | 4 | 4.1115 | 3.5 | 4.4 | 3.6 | 4.2 | 6.3 | 4.3 | 3.6 | < 10 | 4 | 3.8 | < 20 | 3.7 | < 5 | 4.1 | 5 | < 5 | 3 | < 0.1 | 3 | 4 | 3.6 | 4.5 | 3 | 4.1 | 4 | 4.3 |  |  | 51
| 1 | 6 | 4.3 | 4.2158 | 3.9 | 4.1 | 3.7 | 4 | 5.9 | 4.2 | 3.4 | < 10 | 4 | 4.1 | < 20 | 3.8 | < 5 | 4.2 | < 5 | < 5 | 2 | < 0.1 | 3 | 4 | 3.8 | 3.8 | 4 | 4.1 | 3.6 | 4.4
| Mean |  | 4.0667 | 4.2056 | 3.7667 | 4.2167 | 3.6667 | 4.2167 | 6.25 | 4.2167 | 3.5167 |  | 3.8333 | 4 |  | 3.9 |  | 4.15 | 5 |  | 3.5 | 5.2 | 3.1667 | 4 | 3.85 | 4.05 | 2.8 | 3.9667 | 3.8833 | 4.2667
| Median |  | 4.05 | 4.1933 | 3.85 | 4.2 | 3.7 | 4.1 | 6.3 | 4.2 | 3.5 |  | 4 | 4.05 |  | 3.8 |  | 4.15 | 5 |  | 3.5 | 5.2 | 3 | 4 | 3.85 | 4 | 3 | 3.95 | 3.9 | 4.25
| Std Dev. |  | 0.13663 | 0.073533 | 0.21602 | 0.2137 | 0.05164 | 0.40208 | 0.18708 | 0.075277 | 0.13292 |  | 0.40825 | 0.16733 |  | 0.21909 |  | 0.10488 | 0 |  | 1.5166 |  | 0.40825 | 0 | 0.32094 | 0.2429 | 0.83666 | 0.12111 | 0.1472 | 0.12111
| Rel.Std.Dev. |  | 3.36% | 1.748% | 5.735% | 5.068% | 1.408% | 9.535% | 2.993% | 1.785% | 3.78% |  | 10.65% | 4.183% |  | 5.618% |  | 2.527% | 0% |  | 43.33% |  | 12.89% | 0% | 8.336% | 5.998% | 29.88% | 3.053% | 3.79% | 2.838%
| PDM3 |  | 2.791% | 6.302% | -4.792% | 6.582% | -7.32% | 6.582% | 57.98% | 6.582% | -11.11% |  | -3.107% | 1.106% |  | -1.422% |  | 4.897% | 26.38% |  | -11.53% | 31.44% | -19.96% | 1.106% | -2.686% | 2.369% | -29.23% | 0.2631% | -1.843% | 7.846%
| Z-Score (Absolute) |  | 0.26 | 0.68 | 0.67 | 0.72 | 0.97 | 0.24 | 7.18 | 0.72 | 1.43 | NA | 0.46 | 0.05 | NA | 0.5 | NA | 0.51 | 3.12 | NA | 0.97 | 3.74 | 2.51 | 0.05 | 0.41 | 0.07 | 3.64 | 0.05 | 0.13 | 0.87
| Analytical results for Sr in OREAS 905 (Certified Value 157 ppm) | [31 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 151.5 | 161.0972 | 161.2 | 157 | 161 | 159 | 155 | 160 | 158 | 150 | 150 | 141.9 | 130 | 151 | 135 | 162.51 | 156 | 160.6164 | 148 | 155 | 154 | 161 | 157 | 154.5 | 157 | 154.5 | 146 | 164 |  |  | 1
| 1 | 2 | 153 | 163.2709 | 155.6 | 165 | 160.5 | 160.5 | 160 | 170 | 162 | 172 | 149 | 139.9 | 145 | 150 | 144 | 158.31 | 148 | 160.7598 | 150 | 159 | 162 | 160 | 152.5 | 157 | 162 | 152.5 | 147 | 168 |  |  | 0
| 1 | 3 | 162.5 | 162.7641 | 154.6 | 163 | 156 | 162 | 157 | 169 | 155 | 158 | 150 | 166.8 | 123 | 152 | 131 | 160.85 | 147 | 158.4054 | 149 | 155 | 157 | 157 | 157.5 | 159.5 | 162 | 155.5 | 141 | 168 |  |  | 16
| 1 | 4 | 155.5 | 162.4035 | 156.1 | 163 | 159.5 | 159.5 | 153 | 169 | 150 | 155 | 149 | 158.2 | 145 | 151 | 108 | 158.05 | 148 | 156.722 | 149 | 154 | 155 | 165 | 159 | 155.5 | 167 | 155 | 150 | 163 |  |  | 156.9475
| 1 | 5 | 154.5 | 160.8066 | 154.8 | 168 | 159 | 163 | 158 | 176 | 154 | 154 | 134 | 149 | 140 | 153 | 136 | 161.24 | 154 | 157.087 | 148 | 161 | 153 | 158 | 159.5 | 156 | 163 | 156 | 144 | 167 |  |  | 52
| 1 | 6 | 155.5 | 163.0662 | 158.3 | 169 | 160.5 | 162 | 160 | 173 | 160 | 157 | 143 | 161.2 | 144 | 148 | 127 | 161.42 | 148 | 157.8276 | 147 | 156 | 153 | 163 | 157 | 154 | 164 | 160.5 | 141 | 167
| Mean |  | 155.4167 | 162.2347 | 156.7667 | 164.1667 | 159.4167 | 161 | 157.1667 | 169.5 | 156.5 | 157.6667 | 145.8333 | 152.8333 | 137.8333 | 150.8333 | 130.1667 | 160.3967 | 150.1667 | 158.5697 | 148.5 | 156.6667 | 155.6667 | 160.6667 | 157.0833 | 156.0833 | 162.5 | 155.6667 | 144.8333 | 166.1667
| Median |  | 155 | 162.5838 | 155.85 | 164 | 160 | 161.25 | 157.5 | 169.5 | 156.5 | 156 | 149 | 153.6 | 142 | 151 | 133 | 161.045 | 148 | 158.1165 | 148.5 | 155.5 | 154.5 | 160.5 | 157.25 | 155.75 | 162.5 | 155.25 | 145 | 167
| Std Dev. |  | 3.8002 | 1.0399 | 2.5445 | 4.3089 | 1.828 | 1.5811 | 2.7869 | 5.3944 | 4.3704 | 7.5542 | 6.3692 | 10.9121 | 9.2394 | 1.7224 | 12.2543 | 1.8051 | 3.8166 | 1.7421 | 1.0488 | 2.7325 | 3.4448 | 3.0111 | 2.4782 | 1.9854 | 3.2711 | 2.6583 | 3.5449 | 2.137
| Rel.Std.Dev. |  | 2.445% | 0.641% | 1.623% | 2.625% | 1.147% | 0.9821% | 1.773% | 3.183% | 2.793% | 4.791% | 4.367% | 7.14% | 6.703% | 1.142% | 9.414% | 1.125% | 2.542% | 1.099% | 0.7063% | 1.744% | 2.213% | 1.874% | 1.578% | 1.272% | 2.013% | 1.708% | 2.448% | 1.286%
| PDM3 |  | -0.9754% | 3.369% | -0.1152% | 4.6% | 1.573% | 2.582% | 0.1397% | 7.998% | -0.2851% | 0.4582% | -7.081% | -2.621% | -12.18% | -3.896% | -17.06% | 2.198% | -4.32% | 1.034% | -5.382% | -0.1789% | -0.8161% | 2.37% | 0.08656% | -0.5506% | 3.538% | -0.8161% | -7.719% | 5.874%
| Z-Score (Absolute) |  | 0.2 | 0.97 | 0.03 | 1.3 | 0.48 | 0.75 | 0.1 | 2.21 | 0.01 | 0.3 | 1.43 | 0.64 | 3.2 | 0.98 | 3.75 | 0.65 | 1.3 | 0.34 | 1.38 | 0.01 | 0.16 | 0.7 | 0.09 | 0.09 | 1.01 | 0.16 | 2.01 | 1.64
| Analytical results for Ta in OREAS 905 (Certified Value 1.34 ppm) | [31 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 17 | 19 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [7 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 1.28 | 1.3906 | 1.22 | 1.49 | 1.29 | 1.23 | 0.81 | 1.3 | 0.91 | 0.2 | 1.55 | 1.46 | 1.42 | < 5 | 1.5215 | 3.9 | 1.5 | 1.36 | 1.29 | 1.15 | 1.33 | 1.41 | 1.2 | [7 empty] | 1
| 1 | 2 | 1.27 | 1.4211 | 1.37 | 1.48 | 1.3 | 1.23 | 0.8 | 1.4 | 0.61 | 0.9 | 1.66 | 1.34 | 1.42 | < 5 | 1.5865 | 3 | 1.4 | 1.18 | 1.31 | 0.88 | 1.26 | 1.4 | 1.3 | [7 empty] | 14
| 1 | 3 | 1.33 | 1.4188 | 1.45 | 1.47 | 1.27 | 1.27 | 0.76 | 1.2 | 0.58 | 0.9 | 1.51 | 1.25 | 1.35 | < 5 | 2.3485 | 2.7 | 1.4 | 1.44 | 1.34 | 0.94 | 1.3 | 1.29 | 1.3 | [7 empty] | 16
| 1 | 4 | 1.28 | 1.4196 | 1.21 | 1.45 | 1.31 | 1.24 | 0.79 | 1.3 | 0.45 | 1 | 1.38 | 1.3 | 1.38 | < 5 | 2.2555 | 2.2 | 1.4 | 1.41 | 1.33 | 0.9 | 1.29 | 1.44 | 1.2 | [7 empty] | 1.345
| 1 | 5 | 1.3 | 1.4045 | 1.31 | 1.49 | 1.29 | 1.28 | 0.72 | 1.3 | 0.66 | 1 | 1.36 | 1.32 | 1.38 | < 5 | 1.8905 | 2.3 | 1.5 | 1.3 | 1.3 | 0.82 | 1.31 | 1.44 | 1.3 | [7 empty] | 53
| 1 | 6 | 1.32 | 1.3956 | 1.23 | 1.53 | 1.3 | 1.28 | 0.71 | 1.3 | 0.46 | 1 | 1.46 | 1.32 | 1.33 | < 5 | 1.8085 | 2.2 | 1.4 | 1.36 | 1.3 | 0.76 | 1.32 | 1.36 | 1.3
| Mean |  | 1.2967 | 1.4084 | 1.2983 | 1.485 | 1.2933 | 1.255 | 0.765 | 1.3 | 0.61167 | 0.83333 | 1.4867 | 1.3317 | 1.38 |  | 1.9018 | 2.7167 | 1.4333 | 1.3417 | 1.3117 | 0.90833 | 1.3017 | 1.39 | 1.2667
| Median |  | 1.29 | 1.4117 | 1.27 | 1.485 | 1.295 | 1.255 | 0.775 | 1.3 | 0.595 | 0.95 | 1.485 | 1.32 | 1.38 |  | 1.8495 | 2.5 | 1.4 | 1.36 | 1.305 | 0.89 | 1.305 | 1.405 | 1.3
| Std Dev. |  | 0.024221 | 0.013353 | 0.096833 | 0.026646 | 0.013663 | 0.02429 | 0.042308 | 0.063246 | 0.16822 | 0.31411 | 0.11201 | 0.069976 | 0.036332 |  | 0.33987 | 0.66156 | 0.05164 | 0.09261 | 0.019408 | 0.13423 | 0.024833 | 0.057271 | 0.05164
| Rel.Std.Dev. |  | 1.868% | 0.9481% | 7.458% | 1.794% | 1.056% | 1.935% | 5.531% | 4.865% | 27.5% | 37.69% | 7.534% | 5.255% | 2.633% |  | 17.87% | 24.35% | 3.603% | 6.903% | 1.48% | 14.78% | 1.908% | 4.12% | 4.077%
| PDM3 |  | -3.592% | 4.714% | -3.468% | 10.41% | -3.84% | -6.69% | -43.12% | -3.344% | -54.52% | -38.04% | 10.53% | -0.99% | 2.604% |  | 41.4% | 102% | 6.569% | -0.2465% | -2.477% | -32.47% | -3.221% | 3.347% | -5.823%
| Z-Score (Absolute) |  | 0.06 | 0.87 | 0.05 | 1.51 | 0.09 | 0.41 | 4.48 | 0.03 | 5.75 | 2.86 | 1.23 | 0.02 | 0.63 | NA | 4.97 | 11.74 | 1.08 | 0.31 | 0.07 | 3.29 | 0.02 | 0.72 | 0.31
| Analytical results for Tb in OREAS 905 (Certified Value 0.77 ppm) | [31 empty] | Certified
Tb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  | 1
| No | No | 01: 04; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 08: 15; 09: 17; 10: 26; 11: 30; 12: 32; 13: 33 |  |  | ppm
|  |  | 01: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*OES/MS |  |  | 2
3
| 1 | 1 | 01: 0.68; 03: 0.77; 04: 0.76; 05: 0.81; 06: 0.7; 07: 0.7; 08: 0.85; 09: 0.87; 10: 0.82; 11: 1.44; 12: 0.86; 13: 0.8 |  |  | 1
| 1 | 2 | 01: 0.71; 03: 0.8; 04: 0.88; 05: 0.82; 06: 0.7; 07: 0.7; 08: 0.83; 09: 0.83; 10: 0.78; 11: 1.49; 12: 0.87; 13: 0.7 |  |  | 0
| 1 | 3 | 01: 0.71; 03: 0.79; 04: 0.78; 05: 0.81; 06: 0.7; 07: 0.7; 08: 0.78; 09: 0.85; 10: 0.78; 11: 1.48; 12: 0.86; 13: 0.6 |  |  | 16
| 1 | 4 | 01: 0.62; 03: 0.82; 04: 0.76; 05: 0.78; 06: 0.8; 07: 0.7; 08: 0.8; 09: 0.83; 10: 0.78; 11: 1.54; 12: 0.87; 13: 0.7 |  |  | 0.7703
| 1 | 5 | 01: 0.67; 03: 0.82; 04: 0.84; 05: 0.8; 06: 0.6; 07: 0.7; 08: 0.8; 09: 0.86; 10: 0.76; 11: 1.43; 12: 0.86; 13: 0.7 |  |  | 54
| 1 | 6 | 01: 0.65; 03: 0.8; 04: 0.76; 05: 0.82; 06: 0.7; 07: 0.8; 08: 0.8; 09: 0.83; 10: 0.76; 11: 1.47; 12: 0.85; 13: 0.7
| Mean |  | 01: 0.67333; 03: 0.8; 04: 0.79667; 05: 0.80667; 06: 0.7; 07: 0.71667; 08: 0.81; 09: 0.845; 10: 0.78; 11: 1.475; 12: 0.86167; 13: 0.7
| Median |  | 01: 0.675; 03: 0.8; 04: 0.77; 05: 0.81; 06: 0.7; 07: 0.7; 08: 0.8; 09: 0.84; 10: 0.78; 11: 1.475; 12: 0.86; 13: 0.7
| Std Dev. |  | 01: 0.035024; 03: 0.018974; 04: 0.051251; 05: 0.015055; 06: 0.063246; 07: 0.040825; 08: 0.025298; 09: 0.017607; 10: 0.021909; 11: 0.03937; 12: 0.0075277; 13: 0.063246
| Rel.Std.Dev. |  | 01: 5.202%; 03: 2.372%; 04: 6.433%; 05: 1.866%; 06: 9.035%; 07: 5.696%; 08: 3.123%; 09: 2.084%; 10: 2.809%; 11: 2.669%; 12: 0.8736%; 13: 9.035%
| PDM3 |  | 01: -12.59%; 03: 3.855%; 04: 3.423%; 05: 4.721%; 06: -9.127%; 07: -6.963%; 08: 5.153%; 09: 9.697%; 10: 1.259%; 11: 91.48%; 12: 11.86%; 13: -9.127%
| Z-Score (Absolute) |  | 01: 1.24; 03: 0.11; 04: 0.11; 05: 0.18; 06: 0.96; 07: 0.78; 08: 0.21; 09: 0.59; 10: 0.11; 11: 7.29; 12: 0.76; 13: 0.96
| Analytical results for Te in OREAS 905 (Certified Value 0.088 ppm) | [31 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 13 | 15 | 16 | 17 | 20 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [8 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [8 empty] | 3
3
| 1 | 1 | 0.06 | < 0.1 | < 0.05 | < 0.1 | 0.09 | 0.09 | < 0.1 | < 0.2 | 0.11 | < 5 | 0.1 | 0.06 | < 1 | < 0.05 | < 0.2 | < 0.2 | 0.09 | 0.07 | 0.1 | 0.09 | 0.09 | < 0.05 | [8 empty] | 1
| 1 | 2 | 0.08 | < 0.1 | < 0.05 | < 0.1 | 0.08 | 0.08 | < 0.1 | < 0.2 | 0.06 | < 5 | < 0.1 | 0.07 | < 1 | 0.1 | < 0.2 | < 0.2 | 0.08 | 0.08 | 0.2 | 0.09 | 0.08 | < 0.05 | [8 empty] | 8
| 1 | 3 | 0.08 | < 0.1 | < 0.05 | < 0.1 | 0.1 | 0.09 | < 0.1 | < 0.2 | 0.06 | < 5 | < 0.1 | 0.08 | < 1 | 0.11 | < 0.2 | < 0.2 | 0.13 | 0.07 | 0.2 | 0.06 | 0.12 | < 0.05 | [8 empty] | 16
| 1 | 4 | 0.05 | < 0.1 | < 0.05 | < 0.1 | 0.09 | 0.09 | < 0.1 | < 0.2 | < 0.05 | < 5 | < 0.1 | 0.08 | < 1 | 0.1 | < 0.2 | < 0.2 | 0.1 | 0.08 | 0.2 | 0.11 | 0.1 | < 0.05 | [8 empty] | 0.0878
| 1 | 5 | 0.06 | < 0.1 | < 0.05 | < 0.1 | 0.08 | 0.09 | < 0.1 | < 0.2 | < 0.05 | < 5 | < 0.1 | 0.08 | < 1 | 0.07 | 0.24 | < 0.2 | 0.08 | 0.07 | 0.2 | 0.09 | 0.1 | < 0.05 | [8 empty] | 55
| 1 | 6 | 0.05 | < 0.1 | < 0.05 | < 0.1 | 0.08 | 0.09 | < 0.1 | < 0.2 | 0.05 | < 5 | < 0.1 | 0.08 | < 1 | 0.11 | < 0.2 | < 0.2 | 0.1 | 0.09 | 0.2 | 0.12 | 0.11 | < 0.05
| Mean |  | 01: 0.063333; 06: 0.086667; 07: 0.088333; 10: 0.07; 13: 0.1; 15: 0.075; 17: 0.098; 20: 0.24; 28: 0.096667; 29: 0.076667; 30: 0.18333; 31: 0.093333; 32: 0.1
| Median |  | 01: 0.06; 06: 0.085; 07: 0.09; 10: 0.06; 13: 0.1; 15: 0.08; 17: 0.1; 20: 0.24; 28: 0.095; 29: 0.075; 30: 0.2; 31: 0.09; 32: 0.1
| Std Dev. |  | 01: 0.013663; 06: 0.008165; 07: 0.0040825; 10: 0.02708; 15: 0.0083666; 17: 0.016432; 28: 0.018619; 29: 0.008165; 30: 0.040825; 31: 0.020656; 32: 0.014142
| Rel.Std.Dev. |  | 01: 21.57%; 06: 9.421%; 07: 4.622%; 10: 38.69%; 15: 11.16%; 17: 16.77%; 28: 19.26%; 29: 10.65%; 30: 22.27%; 31: 22.13%; 32: 14.14%
| PDM3 |  | 01: -27.87%; 06: -1.291%; 07: 0.6074%; 10: -20.27%; 13: 13.9%; 15: -14.58%; 17: 11.62%; 20: 173.3%; 28: 10.1%; 29: -12.68%; 30: 108.8%; 31: 6.302%; 32: 13.9%
| Z-Score (Absolute) |  | 3.03 | NA | NA | NA | 0.67 | 0.51 | NA | NA | 3.71 | NA | 0.67 | 1.85 | NA | 0.47 | 14.83 | NA | 0.34 | 1.69 | 9.1 | 0 | 0.67 | NA
| Analytical results for Th in OREAS 905 (Certified Value 14.6 ppm) | [31 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 1
2
| 1 | 1 | 14 | 14.4459 | 14.6 | 13.9 | 15.1 | 14.5 | 13.6 | 14.9 | 14.1 | 12.9 | 13.86 | < 20 | 15.2 | 15.03 | 15.311 | 22 | 15.9 | 11 | 16 | 16.3 | 13.3 | 14.4 | 15.3 | 14.3 | 14.6 | [5 empty] | 1
| 1 | 2 | 14 | 14.4877 | 15 | 14.2 | 14.8 | 14.6 | 13.9 | 15 | 14.1 | 12.2 | 13.35 | < 20 | 15 | 15.38 | 15.16 | 21 | 15.6 | 13 | 15.6 | 14.8 | 13.2 | 15 | 14.5 | 14.4 | 16 | [5 empty] | 0
| 1 | 3 | 14.3 | 14.481 | 13.7 | 14.3 | 14.5 | 14.7 | 13.9 | 14.9 | 13.7 | 11.6 | 16.21 | < 20 | 14.4 | 14.82 | 14.911 | 20 | 14.7 | 13 | 15.6 | 15.8 | 13.2 | 13.6 | 15 | 14.1 | 16.1 | [5 empty] | 16
| 1 | 4 | 13.8 | 14.5195 | 13.6 | 14.6 | 14.6 | 14.1 | 13.9 | 14.6 | 13.3 | 12.1 | 15.88 | < 20 | 14.7 | 14.95 | 15.391 | 20 | 14.6 | 11 | 15.3 | 15.4 | 13.3 | 14.3 | 15 | 14.8 | 15.8 | [5 empty] | 14.6454
| 1 | 5 | 14.1 | 14.5132 | 14.3 | 14.5 | 14.8 | 14.7 | 13.6 | 15.1 | 13.8 | 7.6 | 15.36 | < 20 | 14.7 | 15.11 | 14.453 | 20 | 15.5 | 13 | 16.1 | 15.7 | 13.3 | 13.4 | 15.1 | 14.4 | 15.6 | [5 empty] | 56
| 1 | 6 | 13.8 | 14.4671 | 13.6 | 14.4 | 14.5 | 14.9 | 14.1 | 14.7 | 13.7 | 9.9 | 16.39 | < 20 | 14.9 | 15.01 | 14.326 | 20 | 14.4 | 11 | 15.2 | 16 | 13.2 | 14 | 15.6 | 14.4 | 15.5
| Mean |  | 14 | 14.4857 | 14.1333 | 14.3167 | 14.7167 | 14.5833 | 13.8333 | 14.8667 | 13.7833 | 11.05 | 15.175 |  | 14.8167 | 15.05 | 14.9253 | 20.5 | 15.1167 | 12 | 15.6333 | 15.6667 | 13.25 | 14.1167 | 15.0833 | 14.4 | 15.6
| Median |  | 14 | 14.4844 | 14 | 14.35 | 14.7 | 14.65 | 13.9 | 14.9 | 13.75 | 11.85 | 15.62 |  | 14.8 | 15.02 | 15.0355 | 20 | 15.1 | 12 | 15.6 | 15.75 | 13.25 | 14.15 | 15.05 | 14.4 | 15.7
| Std Dev. |  | 0.18974 | 0.027762 | 0.59217 | 0.24833 | 0.23166 | 0.27142 | 0.19664 | 0.18619 | 0.29944 | 1.9685 | 1.2758 |  | 0.27869 | 0.18836 | 0.44792 | 0.83666 | 0.62423 | 1.0954 | 0.36148 | 0.52026 | 0.054772 | 0.58109 | 0.3656 | 0.22804 | 0.54037
| Rel.Std.Dev. |  | 1.355% | 0.1917% | 4.19% | 1.735% | 1.574% | 1.861% | 1.421% | 1.252% | 2.173% | 17.81% | 8.407% |  | 1.881% | 1.252% | 3.001% | 4.081% | 4.129% | 9.129% | 2.312% | 3.321% | 0.4134% | 4.116% | 2.424% | 1.584% | 3.464%
| PDM3 |  | -4.407% | -1.09% | -3.496% | -2.244% | 0.4868% | -0.4236% | -5.545% | 1.511% | -5.886% | -24.55% | 3.616% |  | 1.17% | 2.763% | 1.912% | 39.98% | 3.218% | -18.06% | 6.746% | 6.974% | -9.528% | -3.61% | 2.99% | -1.675% | 6.518%
| Z-Score (Absolute) |  | 0.84 | 0.21 | 0.67 | 0.43 | 0.09 | 0.09 | 1.06 | 0.28 | 1.12 | 3.77 | 0.68 | NA | 0.22 | 0.52 | 0.36 | 7.57 | 0.6 | 3.43 | 1.27 | 1.32 | 1.81 | 0.69 | 0.56 | 0.32 | 1.23
| Analytical results for Ti in OREAS 905 (Certified Value 0.122 wt.%) | [31 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  | wt.%
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  | 3
3
| 1 | 1 | 0.114 | 0.11532 | 0.11 | 0.13 | 0.125 | 0.121 | 0.119 | 0.13 | 0.12 | 0.12 | 0.13 | 0.1155 | 0.13 | 0.12 | 0.12 | 0.1219 | 0.12 | 0.112 | 0.12 | 0.127 | 0.13 | 0.121 | 0.12 | 0.11 | 0.122 | 0.12 | 0.129 |  |  |  | 1
| 1 | 2 | 0.122 | 0.11405 | 0.12 | 0.132 | 0.123 | 0.12 | 0.12 | 0.13 | 0.12 | 0.13 | 0.13 | 0.1143 | 0.13 | 0.11 | 0.12 | 0.122 | 0.12 | 0.112 | 0.13 | 0.128 | 0.13 | 0.121 | 0.121 | 0.096 | 0.121 | 0.12 | 0.129 |  |  |  | 0
| 1 | 3 | 0.127 | 0.11376 | 0.13 | 0.127 | 0.122 | 0.121 | 0.119 | 0.13 | 0.12 | 0.12 | 0.14 | 0.1227 | 0.12 | 0.12 | 0.12 | 0.1191 | 0.11 | 0.112 | 0.12 | 0.127 | 0.13 | 0.12 | 0.124 | 0.102 | 0.122 | 0.12 | 0.131 |  |  |  | 16
| 1 | 4 | 0.124 | 0.11456 | 0.12 | 0.129 | 0.124 | 0.12 | 0.116 | 0.13 | 0.11 | 0.12 | 0.13 | 0.1222 | 0.12 | 0.12 | 0.12 | 0.1208 | 0.12 | 0.112 | 0.12 | 0.127 | 0.13 | 0.118 | 0.122 | 0.098 | 0.121 | 0.12 | 0.13 |  |  |  | 0.12192
| 1 | 5 | 0.122 | 0.1139 | 0.13 | 0.13 | 0.122 | 0.122 | 0.121 | 0.135 | 0.11 | 0.12 | 0.13 | 0.1148 | 0.13 | 0.12 | 0.11 | 0.121 | 0.11 | 0.112 | 0.12 | 0.128 | 0.12 | 0.122 | 0.12 | 0.091 | 0.123 | 0.12 | 0.132 |  |  |  | 57
| 1 | 6 | 0.12 | 0.11461 | 0.12 | 0.132 | 0.125 | 0.121 | 0.117 | 0.135 | 0.12 | 0.12 | 0.13 | 0.1248 | 0.13 | 0.11 | 0.11 | 0.1199 | 0.11 | 0.112 | 0.12 | 0.126 | 0.13 | 0.119 | 0.119 | 0.088 | 0.125 | 0.12 | 0.13
| Mean |  | 0.1215 | 0.11437 | 0.12167 | 0.13 | 0.1235 | 0.12083 | 0.11867 | 0.13167 | 0.11667 | 0.12167 | 0.13167 | 0.11905 | 0.12667 | 0.11667 | 0.11667 | 0.12078 | 0.115 | 0.112 | 0.12167 | 0.12717 | 0.12833 | 0.12017 | 0.121 | 0.0975 | 0.12233 | 0.12 | 0.13017
| Median |  | 0.122 | 0.1143 | 0.12 | 0.13 | 0.1235 | 0.121 | 0.119 | 0.13 | 0.12 | 0.12 | 0.13 | 0.11885 | 0.13 | 0.12 | 0.12 | 0.1209 | 0.115 | 0.112 | 0.12 | 0.127 | 0.13 | 0.1205 | 0.1205 | 0.097 | 0.122 | 0.12 | 0.13
| Std Dev. |  | 0.0043704 | 0.0005821 | 0.0075277 | 0.0018974 | 0.0013784 | 0.00075277 | 0.0018619 | 0.002582 | 0.005164 | 0.0040825 | 0.0040825 | 0.0046805 | 0.005164 | 0.005164 | 0.005164 | 0.0011303 | 0.0054772 | 1.5202e-17 | 0.0040825 | 0.00075277 | 0.0040825 | 0.001472 | 0.0017889 | 0.007893 | 0.0015055 | 0 | 0.001169
| Rel.Std.Dev. |  | 3.597% | 0.509% | 6.187% | 1.46% | 1.116% | 0.623% | 1.569% | 1.961% | 4.426% | 3.355% | 3.101% | 3.932% | 4.077% | 4.426% | 4.426% | 0.9358% | 4.763% | 1.357e-14% | 3.355% | 0.592% | 3.181% | 1.225% | 1.478% | 8.095% | 1.231% | 0% | 0.8981%
| PDM3 |  | -0.3428% | -6.194% | -0.2061% | 6.629% | 1.298% | -0.8896% | -2.667% | 7.996% | -4.307% | -0.2061% | 7.996% | -2.352% | 3.895% | -4.307% | -4.307% | -0.9306% | -5.674% | -8.135% | -0.2061% | 4.305% | 5.262% | -1.436% | -0.7529% | -20.03% | 0.3407% | -1.573% | 6.766%
| Z-Score (Absolute) |  | 0.08 | 1.03 | 0.1 | 1.4 | 0.39 | 0.03 | 0.36 | 1.66 | 0.67 | 0.1 | 1.66 | 0.3 | 0.88 | 0.67 | 0.67 | 0.03 | 0.93 | 1.4 | 0.1 | 0.96 | 1.14 | 0.13 | 0 | 3.66 | 0.21 | 0.16 | 1.43
| Analytical results for Tl in OREAS 905 (Certified Value 0.72 ppm) | [31 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 20 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 0.65 | 0.75942 | 0.71 | 0.7 | 0.73 | 0.73 | 0.64 | 0.7 | 0.76 | < 5 | 0.74 | 0.73 | < 20 | 0.74 | 0.74 | < 10 | 0.755 | 2.8 | 0.9 | 0.8 | 0.67 | 0.6 | 0.71 | 0.74 | 0.76 | [5 empty] | 1
| 1 | 2 | 0.64 | 0.79623 | 0.76 | 0.7 | 0.7 | 0.76 | 0.64 | 0.8 | 0.74 | < 5 | 0.66 | 0.74 | < 20 | 0.77 | 0.71 | < 10 | 0.799 | 2 | 0.8 | 0.66 | 0.68 | 0.6 | 0.67 | 0.76 | 0.8 | [5 empty] | 0
| 1 | 3 | 0.68 | 0.78164 | 0.73 | 0.7 | 0.71 | 0.74 | 0.64 | 0.7 | 0.73 | < 5 | 0.66 | 0.77 | < 20 | 0.74 | 0.71 | < 10 | 0.788 | 1.5 | 0.8 | 0.85 | 0.69 | 0.6 | 0.71 | 0.76 | 0.8 | [5 empty] | 16
| 1 | 4 | 0.65 | 0.74177 | 0.73 | 0.7 | 0.72 | 0.7 | 0.65 | 0.7 | 0.71 | < 5 | 0.64 | 0.75 | < 20 | 0.74 | 0.72 | < 10 | 0.776 | 1.3 | 0.8 | 0.8 | 0.7 | 0.6 | 0.7 | 0.77 | 0.79 | [5 empty] | 0.72195
| 1 | 5 | 0.66 | 0.77779 | 0.74 | 0.7 | 0.74 | 0.72 | 0.65 | 0.7 | 0.74 | < 5 | 0.68 | 0.73 | < 20 | 0.78 | 0.7 | < 10 | 0.805 | 1.3 | 0.8 | 0.72 | 0.71 | 0.6 | 0.71 | 0.75 | 0.82 | [5 empty] | 58
| 1 | 6 | 0.65 | 0.73495 | 0.72 | 0.7 | 0.71 | 0.76 | 0.68 | 0.7 | 0.73 | < 5 | 0.64 | 0.8 | < 20 | 0.75 | 0.7 | < 10 | 0.723 | 1.2 | 0.8 | 0.77 | 0.68 | 0.6 | 0.71 | 0.73 | 0.83
| Mean |  | 0.655 | 0.7653 | 0.73167 | 0.7 | 0.71833 | 0.735 | 0.65 | 0.71667 | 0.735 |  | 0.67 | 0.75333 |  | 0.75333 | 0.71333 |  | 0.77433 | 1.6833 | 0.81667 | 0.76667 | 0.68833 | 0.6 | 0.70167 | 0.75167 | 0.8
| Median |  | 0.65 | 0.76861 | 0.73 | 0.7 | 0.715 | 0.735 | 0.645 | 0.7 | 0.735 |  | 0.66 | 0.745 |  | 0.745 | 0.71 |  | 0.782 | 1.4 | 0.8 | 0.785 | 0.685 | 0.6 | 0.71 | 0.755 | 0.8
| Std Dev. |  | 0.013784 | 0.024039 | 0.017224 | 1.2162e-16 | 0.01472 | 0.023452 | 0.015492 | 0.040825 | 0.016432 |  | 0.037417 | 0.027325 |  | 0.017512 | 0.015055 |  | 0.030813 | 0.61779 | 0.040825 | 0.067429 | 0.01472 | 0 | 0.016021 | 0.01472 | 0.024495
| Rel.Std.Dev. |  | 2.104% | 3.141% | 2.354% | 1.737e-14% | 2.049% | 3.191% | 2.383% | 5.696% | 2.236% |  | 5.585% | 3.627% |  | 2.325% | 2.111% |  | 3.979% | 36.7% | 4.999% | 8.795% | 2.138% | 0% | 2.283% | 1.958% | 3.062%
| PDM3 |  | -9.274% | 6.004% | 1.346% | -3.04% | -0.5011% | 1.807% | -9.966% | -0.7319% | 1.807% |  | -7.196% | 4.347% |  | 4.347% | -1.194% |  | 7.256% | 133.2% | 13.12% | 6.194% | -4.656% | -16.89% | -2.81% | 4.116% | 10.81%
| Z-Score (Absolute) |  | 1.72 | 0.66 | 0.03 | 0.69 | 0.31 | 0.03 | 1.84 | 0.34 | 0.03 | NA | 1.6 | 0.41 | NA | 0.41 | 0.41 | NA | 0.85 | 15.01 | 1.38 | 0.69 | 0.93 | 2.75 | 0.52 | 0.38 | 1.38
| Analytical results for Tm in OREAS 905 (Certified Value 0.11 ppm) | [31 empty] | Certified
Tm | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab | [5 empty] | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 17; 08: 26; 09: 30; 10: 33 | [5 empty] | ppm
|  |  | 01: 4A*MS; 03: 4A*MS; 04: 4A*OES/MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*MS; 10: 4A*OES/MS | [5 empty] | 2
2
| 1 | 1 | 01: 0.11; 03: 0.1; 04: 0.12; 05: 0.1; 06: < 0.1; 07: 0.13; 08: 0.12; 09: 0.09; 10: 0.1 | [5 empty] | 1
| 1 | 2 | 01: 0.12; 03: 0.2; 04: 0.12; 05: < 0.1; 06: < 0.1; 07: 0.14; 08: 0.12; 09: 0.09; 10: 0.1 | [5 empty] | 0
| 1 | 3 | 01: 0.11; 03: 0.1; 04: 0.13; 05: 0.1; 06: 0.1; 07: 0.12; 08: 0.12; 09: 0.09; 10: 0.1 | [5 empty] | 16
| 1 | 4 | 01: 0.12; 03: 0.1; 04: 0.12; 05: 0.1; 06: 0.1; 07: 0.13; 08: 0.12; 09: 0.09; 10: 0.1 | [5 empty] | 0.10889
| 1 | 5 | 01: 0.12; 03: 0.1; 04: 0.13; 05: 0.1; 06: < 0.1; 07: 0.12; 08: 0.12; 09: 0.09; 10: 0.1 | [5 empty] | 59
| 1 | 6 | 01: 0.12; 03: 0.1; 04: 0.12; 05: 0.1; 06: 0.1; 07: 0.14; 08: 0.12; 09: 0.09; 10: 0.2
| Mean |  | 01: 0.11667; 03: 0.11667; 04: 0.12333; 05: 0.1; 06: 0.1; 07: 0.13; 08: 0.12; 09: 0.09; 10: 0.11667
| Median |  | 01: 0.12; 03: 0.1; 04: 0.12; 05: 0.1; 06: 0.1; 07: 0.13; 08: 0.12; 09: 0.09; 10: 0.1
| Std Dev. |  | 01: 0.005164; 03: 0.040825; 04: 0.005164; 05: 0; 06: 1.6997e-17; 07: 0.0089443; 08: 0; 09: 1.5202e-17; 10: 0.040825
| Rel.Std.Dev. |  | 01: 4.426%; 03: 34.99%; 04: 4.187%; 05: 0%; 06: 1.7e-14%; 07: 6.88%; 08: 0%; 09: 1.689e-14%; 10: 34.99%
| PDM3 |  | 01: 7.143%; 03: 7.143%; 04: 13.27%; 05: -8.163%; 06: -8.163%; 07: 19.39%; 08: 10.2%; 09: -17.35%; 10: 7.143%
| Z-Score (Absolute) |  | 01: 1.12; 03: 0; 04: 1.57; 05: 0; 06: 0; 07: 2.02; 08: 1.35; 09: 0.67; 10: 0
| Analytical results for U in OREAS 905 (Certified Value 4.97 ppm) | [31 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 4.6 | 5.0461 | 5.4 | 5.1 | 4.9 | 4.4 | 4.79 | 4.8 | 4.8 | < 10 | < 10 | 4.88 | < 20 | 5.24 | 5.11 | 5.124 | < 20 | 5.7 | < 20 | 5.6 | 5.6 | 4.5 | 4.97 | 4.9 | 4.79 | 5.5 |  |  |  |  | 1
| 1 | 2 | 4.6 | 4.8894 | 5.7 | 5.25 | 4.8 | 4.5 | 5.03 | 4.8 | 4.76 | < 10 | < 10 | 4.9 | < 20 | 5.03 | 5.12 | 5.029 | < 20 | 5.3 | < 20 | 5.3 | 4.9 | 4.5 | 5.09 | 4.7 | 4.77 | 5.5 |  |  |  |  | 0
| 1 | 3 | 4.7 | 4.9757 | 5.4 | 5.17 | 4.7 | 4.6 | 4.9 | 4.7 | 4.61 | < 10 | < 10 | 5.26 | < 20 | 4.91 | 4.97 | 5.084 | < 20 | 5.2 | < 20 | 5.4 | 5.2 | 4.5 | 4.78 | 4.8 | 4.81 | 5.6 |  |  |  |  | 16
| 1 | 4 | 4.6 | 4.8731 | 5.5 | 5.3 | 4.8 | 4.4 | 4.85 | 4.7 | 4.52 | < 10 | < 10 | 5.23 | < 20 | 4.96 | 5.02 | 5.071 | < 20 | 5.1 | < 20 | 5.2 | 5 | 4.5 | 5.06 | 4.8 | 4.97 | 5.6 |  |  |  |  | 4.9743
| 1 | 5 | 4.6 | 4.9593 | 5.4 | 5.35 | 4.9 | 4.6 | 4.94 | 5 | 4.79 | < 10 | < 10 | 4.94 | < 20 | 5 | 4.99 | 4.763 | < 20 | 5.2 | < 20 | 5.4 | 5.4 | 4.6 | 4.64 | 4.8 | 4.82 | 5.4 |  |  |  |  | 60
| 1 | 6 | 4.7 | 5.0766 | 5.2 | 5.24 | 4.8 | 4.6 | 4.94 | 4.8 | 4.62 | < 10 | < 10 | 5.34 | < 20 | 4.95 | 5.08 | 4.82 | < 20 | 5.3 | < 20 | 5.2 | 5.4 | 4.5 | 4.74 | 4.9 | 4.75 | 5.7
| Mean |  | 4.6333 | 4.97 | 5.4333 | 5.235 | 4.8167 | 4.5167 | 4.9083 | 4.8 | 4.6833 |  |  | 5.0917 |  | 5.015 | 5.0483 | 4.9818 |  | 5.3 |  | 5.35 | 5.25 | 4.5167 | 4.88 | 4.8167 | 4.8183 | 5.55
| Median |  | 4.6 | 4.9675 | 5.4 | 5.245 | 4.8 | 4.55 | 4.92 | 4.8 | 4.69 |  |  | 5.085 |  | 4.98 | 5.05 | 5.05 |  | 5.25 |  | 5.35 | 5.3 | 4.5 | 4.875 | 4.8 | 4.8 | 5.55
| Std Dev. |  | 0.05164 | 0.081494 | 0.1633 | 0.08961 | 0.075277 | 0.098319 | 0.082805 | 0.10954 | 0.1157 |  |  | 0.20673 |  | 0.11777 | 0.063692 | 0.15159 |  | 0.20976 |  | 0.15166 | 0.26646 | 0.040825 | 0.18536 | 0.075277 | 0.078592 | 0.10488
| Rel.Std.Dev. |  | 1.115% | 1.64% | 3.006% | 1.712% | 1.563% | 2.177% | 1.687% | 2.282% | 2.47% |  |  | 4.06% |  | 2.348% | 1.262% | 3.043% |  | 3.958% |  | 2.835% | 5.075% | 0.9039% | 3.798% | 1.563% | 1.631% | 1.89%
| PDM3 |  | -6.854% | -0.08531% | 9.229% | 5.241% | -3.169% | -9.2% | -1.326% | -3.504% | -5.849% |  |  | 2.36% |  | 0.8186% | 1.489% | 0.1519% |  | 6.548% |  | 7.553% | 5.543% | -9.2% | -1.895% | -3.169% | -3.135% | 11.57%
| Z-Score (Absolute) |  | 1.25 | 0 | 1.72 | 0.98 | 0.57 | 1.68 | 0.23 | 0.63 | 1.06 | NA | NA | 0.45 | NA | 0 | 0.29 | 0.04 | NA | 0.93 | NA | 1.41 | 1.04 | 1.68 | 0.33 | 0.57 | 0.67 | 2.15
| Analytical results for V in OREAS 905 (Certified Value 10.1 ppm) | [31 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 1
2
| 1 | 1 | 9 | 10.5639 | 11 | 9 | 9 | 10 | 10 | 10 | 15 | 13 | 11 | 9 | 11 | 8 | 10 | 8 | 12 | 10.6 | 10.645 | 10 | 11 | 10 | 10 | 10 | 9 | 12 | 10 | 12 | 10 |  | 1
| 1 | 2 | 9 | 10.6156 | 11 | 9 | 9 | 10 | 10 | 10 | 15 | 14 | 11 | 10 | 10 | 8 | 11 | 9 | 12 | 10.6 | 11.166 | 10 | 11 | 10 | 10 | 10 | 9 | 12 | 10 | 11 | 10 |  | 0
| 1 | 3 | 10 | 10.7601 | 11 | 8 | 9 | 10 | 10 | 10 | 14 | 12 | 11 | 9 | 10 | 8 | 10 | 9 | 12 | 10.6 | 10.496 | 10 | 11 | 10 | 10 | 9 | 9 | 12 | 10 | 11 | 10 |  | 16
| 1 | 4 | 10 | 10.5779 | 11 | 9 | 9 | 10 | 10 | 10 | 14 | 13 | 11 | 9 | 10 | 8 | 11 | 10 | 12 | 10.6 | 11.656 | 10 | 11 | 10 | 10 | 10 | 9 | 12 | 10 | 11 | 10 |  | 10.1466
| 1 | 5 | 10 | 10.5181 | 11 | 9 | 9 | 10 | 10 | 10 | 14 | 13 | 11 | 9 | 11 | 8 | 11 | 9 | 12 | 10.6 | 10.557 | 10 | 11 | 10 | 10 | 10 | 9 | 12 | 10 | 11 | 10 |  | 61
| 1 | 6 | 10 | 10.677 | 11 | 9 | 10 | 10 | 10 | 10 | 14 | 13 | 10 | 10 | 11 | 8 | 11 | 9 | 12 | 11.7 | 10.92 | 10 | 12 | 9 | 10 | 10 | 9 | 12 | 10 | 10 | 10
| Mean |  | 9.6667 | 10.6188 | 11 | 8.8333 | 9.1667 | 10 | 10 | 10 | 14.3333 | 13 | 10.8333 | 9.3333 | 10.5 | 8 | 10.6667 | 9 | 12 | 10.7833 | 10.9067 | 10 | 11.1667 | 9.8333 | 10 | 9.8333 | 9 | 12 | 10 | 11 | 10
| Median |  | 10 | 10.5968 | 11 | 9 | 9 | 10 | 10 | 10 | 14 | 13 | 11 | 9 | 10.5 | 8 | 11 | 9 | 12 | 10.6 | 10.7825 | 10 | 11 | 10 | 10 | 10 | 9 | 12 | 10 | 11 | 10
| Std Dev. |  | 0.5164 | 0.087391 | 0 | 0.40825 | 0.40825 | 0 | 0 | 0 | 0.5164 | 0.63246 | 0.40825 | 0.5164 | 0.54772 | 0 | 0.5164 | 0.63246 | 0 | 0.44907 | 0.44459 | 0 | 0.40825 | 0.40825 | 0 | 0.40825 | 0 | 0 | 0 | 0.63246 | 0
| Rel.Std.Dev. |  | 5.342% | 0.823% | 0% | 4.622% | 4.454% | 0% | 0% | 0% | 3.603% | 4.865% | 3.768% | 5.533% | 5.216% | 0% | 4.841% | 7.027% | 0% | 4.165% | 4.076% | 0% | 3.656% | 4.152% | 0% | 4.152% | 0% | 0% | 0% | 5.75% | 0%
| PDM3 |  | -4.73% | 4.653% | 8.41% | -12.94% | -9.658% | -1.445% | -1.445% | -1.445% | 41.26% | 28.12% | 6.768% | -8.015% | 3.483% | -21.16% | 5.125% | -11.3% | 18.27% | 6.275% | 7.491% | -1.445% | 10.05% | -3.088% | -1.445% | -3.088% | -11.3% | 18.27% | -1.445% | 8.41% | -1.445%
| Z-Score (Absolute) |  | 0.34 | 0.63 | 1.01 | 1.18 | 0.84 | 0 | 0 | 0 | 4.38 | 3.03 | 0.84 | 0.67 | 0.51 | 2.02 | 0.67 | 1.01 | 2.02 | 0.61 | 0.92 | 0 | 1.18 | 0.17 | 0 | 0.17 | 1.01 | 2.02 | 0 | 1.01 | 0
| Analytical results for W in OREAS 905 (Certified Value 2.78 ppm) | [31 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 2
3
| 1 | 1 | 2.3 | 2.7151 | 2.5 | 2.7 | 2.7 | 2.9 | 2.7 | 3 | 2.9 | < 10 | < 5 | 2.9 | < 20 | 2.9 | 2.9 | 4 | < 5 | 2.7073 | 4 | 3.4 | 4 | 3.5 | 2.9 | 2.7 | 2.3 | 3 | 2.6 | 2.8 |  |  | 1
| 1 | 2 | 2.3 | 2.8795 | 2.8 | 2.7 | 2.6 | 2.9 | 2.7 | 3 | 2.7 | < 10 | < 5 | 2.9 | < 20 | 2.8 | 3 | 4 | < 5 | 2.9193 | < 4 | 3.1 | 4 | 2.5 | 2.5 | 2.8 | 1.8 | 2.7 | 2.8 | 2.6 |  |  | 0
| 1 | 3 | 2.3 | 2.8832 | 2.8 | 2.7 | 2.5 | 2.9 | 2.6 | 2.5 | 2.6 | < 10 | < 5 | 2.8 | < 20 | 2.7 | 2.9 | 4 | < 5 | 2.9763 | < 4 | 2.4 | 6 | 2.5 | 2.6 | 2.9 | 2 | 2.8 | 2.8 | 2.9 |  |  | 16
| 1 | 4 | 2.3 | 2.8679 | 2.8 | 2.7 | 2.6 | 3 | 2.4 | 3 | 2.4 | < 10 | 5 | 2.9 | < 20 | 2.7 | 2.9 | 3 | < 5 | 2.8584 | < 4 | 2.5 | < 4 | 2.5 | 2.6 | 2.9 | 2 | 2.9 | 2.9 | 2.7 |  |  | 2.7768
| 1 | 5 | 2.4 | 2.7275 | 2.7 | 2.7 | 2.6 | 2.9 | 2.9 | 3 | 2.7 | < 10 | < 5 | 2.6 | < 20 | 2.8 | 2.8 | 5 | < 5 | 2.9493 | < 4 | 3 | < 4 | 2.5 | 2.6 | 2.9 | 1.9 | 2.9 | 3 | 2.5 |  |  | 62
| 1 | 6 | 2.3 | 2.7348 | 2.9 | 2.7 | 2.5 | 2.9 | 2.7 | 3 | 2.5 | < 10 | 5 | 3.1 | < 20 | 2.8 | 3 | 2 | < 5 | 3.0784 | < 4 | 3 | < 4 | 2.5 | 2.7 | 3.2 | 1.8 | 3.2 | 2.6 | 2.6
| Mean |  | 2.3167 | 2.8013 | 2.75 | 2.7 | 2.5833 | 2.9167 | 2.6667 | 2.9167 | 2.6333 |  | 5 | 2.8667 |  | 2.7833 | 2.9167 | 3.6667 |  | 2.9148 | 4 | 2.9 | 4.6667 | 2.6667 | 2.65 | 2.9 | 1.9667 | 2.9167 | 2.7833 | 2.6833
| Median |  | 2.3 | 2.8014 | 2.8 | 2.7 | 2.6 | 2.9 | 2.7 | 3 | 2.65 |  | 5 | 2.9 |  | 2.8 | 2.9 | 4 |  | 2.9343 | 4 | 3 | 4 | 2.5 | 2.6 | 2.9 | 1.95 | 2.9 | 2.8 | 2.65
| Std Dev. |  | 0.040825 | 0.083133 | 0.13784 | 4.8648e-16 | 0.075277 | 0.040825 | 0.1633 | 0.20412 | 0.17512 |  | 0 | 0.1633 |  | 0.075277 | 0.075277 | 1.0328 |  | 0.12489 |  | 0.37947 | 1.1547 | 0.40825 | 0.13784 | 0.16733 | 0.18619 | 0.17224 | 0.16021 | 0.1472
| Rel.Std.Dev. |  | 1.762% | 2.968% | 5.012% | 1.802e-14% | 2.914% | 1.4% | 6.124% | 6.999% | 6.65% |  | 0% | 5.696% |  | 2.705% | 2.581% | 28.17% |  | 4.285% |  | 13.09% | 24.74% | 15.31% | 5.202% | 5.77% | 9.467% | 5.905% | 5.756% | 5.486%
| PDM3 |  | -16.57% | 0.8831% | -0.9657% | -2.766% | -6.968% | 5.036% | -3.967% | 5.036% | -5.167% |  | 80.06% | 3.236% |  | 0.2347% | 5.036% | 32.05% |  | 4.971% | 44.05% | 4.436% | 68.06% | -3.967% | -4.567% | 4.436% | -29.18% | 5.036% | 0.2347% | -3.367%
| Z-Score (Absolute) |  | 2.79 | 0.01 | 0 | 0.58 | 1.25 | 0.67 | 0.46 | 1.16 | 0.96 | NA | 12.72 | 0.39 | NA | 0.1 | 0.67 | 5.01 | NA | 0.66 | 6.94 | 0 | 10.79 | 1.73 | 1.16 | 0.23 | 4.82 | 0.67 | 0.1 | 0.67
| Analytical results for Y in OREAS 905 (Certified Value 15.7 ppm) | [31 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  |  |  | 1
2
| 1 | 1 | 15.6 | 14.3733 | 15.3 | 16.5 | 16.4 | 16 | 15.5 | 16.1 | 15.6 | 15 | 16 | 13.5 | 9 | 16.2 | 16.23 | 16.1 | 18 | 14.9 | 14 | 19 | 16 | 14.4 | 15.2 | 16.2 | 16.5 | 15.8 |  |  |  |  | 1
| 1 | 2 | 15.8 | 14.754 | 16.2 | 16.2 | 16.2 | 16.2 | 15.2 | 16.3 | 15.7 | 17 | 16 | 13.6 | 11 | 16.4 | 15.76 | 15.4 | 18 | 15.7 | 14 | 18.5 | 14.4 | 15 | 15 | 16.2 | 16.6 | 16.7 |  |  |  |  | 0
| 1 | 3 | 16.2 | 14.6622 | 15.9 | 15.4 | 15.6 | 16.3 | 15 | 16.1 | 15.8 | 15 | 17 | 16 | 9 | 16.5 | 15.75 | 15 | 17 | 14.4 | 15 | 18.4 | 15.5 | 15.1 | 15.5 | 16.1 | 16.6 | 16.7 |  |  |  |  | 16
| 1 | 4 | 15.4 | 14.8121 | 15 | 15.6 | 16.1 | 16.1 | 15.2 | 15.9 | 15.5 | 15 | 17 | 16.1 | 11 | 16.5 | 15.73 | 15.3 | 17 | 14.7 | 15 | 18.5 | 15.8 | 15.2 | 15.6 | 16.3 | 16.7 | 16.1 |  |  |  |  | 15.7134
| 1 | 5 | 15.5 | 14.8404 | 14.9 | 15.6 | 16 | 16.1 | 15.6 | 16.5 | 15.9 | 15 | 14 | 14.9 | 11 | 16.6 | 16.07 | 16.6 | 17 | 15.4 | 15 | 18.1 | 15.7 | 15 | 15.4 | 16.6 | 16.4 | 16.8 |  |  |  |  | 63
| 1 | 6 | 15.5 | 14.3239 | 15.4 | 15.7 | 15.6 | 16.1 | 15.7 | 16.3 | 15.8 | 15 | 15 | 16.3 | 13 | 16.5 | 16.3 | 15.5 | 16 | 14.4 | 14 | 18.6 | 16.3 | 14.4 | 14.2 | 16.8 | 16 | 16.2
| Mean |  | 15.6667 | 14.6277 | 15.45 | 15.8333 | 15.9833 | 16.1333 | 15.3667 | 16.2 | 15.7167 | 15.3333 | 15.8333 | 15.0667 | 10.6667 | 16.45 | 15.9733 | 15.65 | 17.1667 | 14.9167 | 14.5 | 18.5167 | 15.6167 | 14.85 | 15.15 | 16.3667 | 16.4667 | 16.3833
| Median |  | 15.55 | 14.7081 | 15.35 | 15.65 | 16.05 | 16.1 | 15.35 | 16.2 | 15.75 | 15 | 16 | 15.45 | 11 | 16.5 | 15.915 | 15.45 | 17 | 14.8 | 14.5 | 18.5 | 15.75 | 15 | 15.3 | 16.25 | 16.55 | 16.45
| Std Dev. |  | 0.29439 | 0.22512 | 0.50892 | 0.42269 | 0.32506 | 0.10328 | 0.27325 | 0.20976 | 0.1472 | 0.8165 | 1.169 | 1.2723 | 1.5055 | 0.13784 | 0.25944 | 0.58907 | 0.75277 | 0.53448 | 0.54772 | 0.29269 | 0.65549 | 0.35637 | 0.51284 | 0.27325 | 0.25033 | 0.40702
| Rel.Std.Dev. |  | 1.879% | 1.539% | 3.294% | 2.67% | 2.034% | 0.6402% | 1.778% | 1.295% | 0.9366% | 5.325% | 7.383% | 8.444% | 14.11% | 0.8379% | 1.624% | 3.764% | 4.385% | 3.583% | 3.777% | 1.581% | 4.197% | 2.4% | 3.385% | 1.67% | 1.52% | 2.484%
| PDM3 |  | -0.2973% | -6.909% | -1.676% | 0.7634% | 1.718% | 2.673% | -2.206% | 3.097% | 0.02095% | -2.419% | 0.7634% | -4.116% | -32.12% | 4.688% | 1.654% | -0.4033% | 9.249% | -5.07% | -7.722% | 17.84% | -0.6155% | -5.495% | -3.585% | 4.158% | 4.794% | 4.264%
| Z-Score (Absolute) |  | 0.33 | 1.77 | 0.5 | 0.09 | 0.32 | 0.55 | 0.63 | 0.65 | 0.09 | 0.68 | 0.09 | 1.09 | 7.86 | 1.04 | 0.31 | 0.19 | 2.14 | 1.32 | 1.96 | 4.22 | 0.13 | 1.42 | 0.67 | 0.91 | 1.21 | 0.94
| Analytical results for Yb in OREAS 905 (Certified Value 0.68 ppm) | [31 empty] | Certified
Yb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  |  |  |  | 1
| No | No | 01: 04; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 08: 15; 09: 17; 10: 26; 11: 30; 12: 32; 13: 33 |  |  |  |  | ppm
|  |  | 01: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*OES/MS; 13: 4A*OES/MS |  |  |  |  | 2
3
| 1 | 1 | 01: 0.6; 03: 0.7; 04: 0.7; 05: 0.7; 06: 0.6; 07: 0.6; 08: 0.7; 09: 0.69; 10: 0.75; 11: 1; 12: 0.7; 13: 0.8 |  |  |  |  | 1
| 1 | 2 | 01: 0.6; 03: 0.7; 04: 0.75; 05: 0.7; 06: 0.6; 07: 0.6; 08: 0.7; 09: 0.72; 10: 0.75; 11: 1; 12: 0.7; 13: 0.8 |  |  |  |  | 0
| 1 | 3 | 01: 0.6; 03: 0.7; 04: 0.65; 05: 0.7; 06: 0.6; 07: 0.6; 08: 0.7; 09: 0.66; 10: 0.7; 11: 1; 12: 0.7; 13: 0.8 |  |  |  |  | 16
| 1 | 4 | 01: 0.6; 03: 0.7; 04: 0.65; 05: 0.7; 06: 0.6; 07: 0.6; 08: 0.7; 09: 0.68; 10: 0.7; 11: 1; 12: 0.7; 13: 0.9 |  |  |  |  | 0.675
| 1 | 5 | 01: 0.6; 03: 0.7; 04: 0.7; 05: 0.7; 06: 0.5; 07: 0.6; 08: 0.7; 09: 0.74; 10: 0.75; 11: 0.9; 12: 0.7; 13: 0.8 |  |  |  |  | 64
| 1 | 6 | 01: 0.5; 03: 0.7; 04: 0.7; 05: 0.7; 06: 0.7; 07: 0.7; 08: 0.7; 09: 0.69; 10: 0.7; 11: 0.9; 12: 0.7; 13: 0.9
| Mean |  | 01: 0.58333; 03: 0.7; 04: 0.69167; 05: 0.7; 06: 0.6; 07: 0.61667; 08: 0.7; 09: 0.69667; 10: 0.725; 11: 0.96667; 12: 0.7; 13: 0.83333
| Median |  | 01: 0.6; 03: 0.7; 04: 0.7; 05: 0.7; 06: 0.6; 07: 0.6; 08: 0.7; 09: 0.69; 10: 0.725; 11: 1; 12: 0.7; 13: 0.8
| Std Dev. |  | 01: 0.040825; 03: 1.2162e-16; 04: 0.037639; 05: 1.2162e-16; 06: 0.063246; 07: 0.040825; 08: 1.2162e-16; 09: 0.028752; 10: 0.027386; 11: 0.05164; 12: 1.2162e-16; 13: 0.05164
| Rel.Std.Dev. |  | 01: 6.999%; 03: 1.737e-14%; 04: 5.442%; 05: 1.737e-14%; 06: 10.54%; 07: 6.62%; 08: 1.737e-14%; 09: 4.127%; 10: 3.777%; 11: 5.342%; 12: 1.737e-14%; 13: 6.197%
| PDM3 |  | 01: -13.58%; 03: 3.704%; 04: 2.469%; 05: 3.704%; 06: -11.11%; 07: -8.642%; 08: 3.704%; 09: 3.21%; 10: 7.407%; 11: 43.21%; 12: 3.704%; 13: 23.46%
| Z-Score (Absolute) |  | 01: 4.05; 03: 0; 04: 0.34; 05: 0; 06: 3.24; 07: 3.37; 08: 0; 09: 0.13; 10: 1.01; 11: 10.79; 12: 0; 13: 5.39
| Analytical results for Zn in OREAS 905 (Certified Value 138 ppm) | [31 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  | 0
0
| 1 | 1 | 132 | 142.4956 | 129 | 155 | 140 | 135 | 139 | 145 | 136 | 137 | 133 | 146 | 164 | 144 | 144 | 131 | 141 | 133 | 134.7762 | 150 | 143 | 146 | 135 | 137 | 134 | 128 | 136 | 126 | 148.2 |  | 1
| 1 | 2 | 135 | 139.603 | 132 | 156 | 138 | 135 | 142 | 145 | 131 | 143 | 129 | 136 | 155 | 147 | 131 | 131 | 134 | 142 | 134.42 | 143 | 143 | 157 | 140 | 136 | 135 | 126 | 137 | 125 | 148 |  | 0
| 1 | 3 | 146 | 142.892 | 139 | 152 | 135 | 137 | 140 | 145 | 133 | 137 | 135 | 142 | 148 | 145 | 134 | 130 | 145 | 145 | 132.7933 | 139 | 143 | 150 | 140 | 133 | 138 | 130 | 138 | 122 | 148.9 |  | 16
| 1 | 4 | 143 | 141.3238 | 130 | 156 | 137 | 135 | 139 | 145 | 129 | 133 | 133 | 138 | 152 | 147 | 135 | 130 | 131 | 150 | 136.8038 | 139 | 167 | 151 | 140 | 133 | 135 | 129 | 138 | 127 | 147.8 |  | 137.9507
| 1 | 5 | 141 | 138.9939 | 136 | 154 | 139 | 138 | 136 | 150 | 131 | 139 | 131 | 131 | 155 | 151 | 152 | 134 | 132 | 142 | 130.0581 | 134 | 146 | 147 | 135 | 137 | 133 | 131 | 137 | 123 | 153.1 |  | 65
| 1 | 6 | 141 | 144.0704 | 133 | 163 | 141 | 138 | 140 | 150 | 134 | 138 | 134 | 133 | 154 | 139 | 147 | 133 | 141 | 143 | 130.3752 | 138 | 142 | 145 | 140 | 135 | 134 | 126 | 144 | 123 | 156.6
| Mean |  | 139.6667 | 141.5631 | 133.1667 | 156 | 138.3333 | 136.3333 | 139.3333 | 146.6667 | 132.3333 | 137.8333 | 132.5 | 137.6667 | 154.6667 | 145.5 | 140.5 | 131.5 | 137.3333 | 142.5 | 133.2044 | 140.5 | 147.3333 | 149.3333 | 138.3333 | 135.1667 | 134.8333 | 128.3333 | 138.3333 | 124.3333 | 150.4333
| Median |  | 141 | 141.9097 | 132.5 | 155.5 | 138.5 | 136 | 139.5 | 145 | 132 | 137.5 | 133 | 137 | 154.5 | 146 | 139.5 | 131 | 137.5 | 142.5 | 133.6067 | 139 | 143 | 148.5 | 140 | 135.5 | 134.5 | 128.5 | 137.5 | 124 | 148.55
| Std Dev. |  | 5.2026 | 1.9709 | 3.7639 | 3.7417 | 2.1602 | 1.5055 | 1.9664 | 2.582 | 2.5033 | 3.2506 | 2.1679 | 5.6095 | 5.2789 | 3.9875 | 8.3606 | 1.6432 | 5.7504 | 5.5408 | 2.6449 | 5.4681 | 9.7297 | 4.4121 | 2.582 | 1.8348 | 1.7224 | 2.0656 | 2.8752 | 1.9664 | 3.6148
| Rel.Std.Dev. |  | 3.725% | 1.392% | 2.826% | 2.398% | 1.562% | 1.104% | 1.411% | 1.76% | 1.892% | 2.358% | 1.636% | 4.075% | 3.413% | 2.741% | 5.951% | 1.25% | 4.187% | 3.888% | 1.986% | 3.892% | 6.604% | 2.955% | 1.866% | 1.357% | 1.277% | 1.61% | 2.078% | 1.582% | 2.403%
| PDM3 |  | 1.244% | 2.619% | -3.468% | 13.08% | 0.2774% | -1.172% | 1.002% | 6.318% | -4.072% | -0.08504% | -3.951% | -0.2059% | 12.12% | 5.472% | 1.848% | -4.676% | -0.4475% | 3.298% | -3.441% | 1.848% | 6.801% | 8.251% | 0.2774% | -2.018% | -2.26% | -6.972% | 0.2774% | -9.871% | 9.049%
| Z-Score (Absolute) |  | 0.22 | 0.52 | 0.84 | 2.63 | 0 | 0.32 | 0.16 | 1.35 | 0.97 | 0.08 | 0.94 | 0.11 | 2.64 | 1.16 | 0.35 | 1.11 | 0.16 | 0.67 | 0.83 | 0.04 | 0.82 | 1.78 | 0 | 0.51 | 0.57 | 1.62 | 0.18 | 2.27 | 1.76
| Analytical results for Zr in OREAS 905 (Certified Value 252 ppm) | [31 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES/MS | 4A*OES | 4A*OES/MS | 4A*OES/MS | 4A*OES/MS |  |  | 0
0
| 1 | 1 | 243 | 267.0296 | 257.7 | 252 | 257 | 245 | 270 | 255 | 260 | 248 | 193 | 236.8 | 258 | 247 | 240.4 | 277.9 | 240 | 231 | 294 | 262.6 | 243 | 297 | 244 | 255 | 255 | 247 | 264 | 247.5 |  |  | 1
| 1 | 2 | 248 | 263.6161 | 267.8 | 242 | 253 | 251 | 278 | 260 | 249 | 268 | 260 | 242.1 | 264 | 243 | 241.9 | 268.9 | 237 | 234 | 288 | 268.5 | 253 | 283 | 232 | 258 | 247 | 248 | 249 | 256.9 |  |  | 0
| 1 | 3 | 260 | 264.7438 | 261.8 | 242 | 249 | 249 | 274 | 260 | 248 | 254 | 258 | 268 | 255 | 244 | 247.2 | 269.4 | 233 | 232 | 274 | 264.6 | 238 | 277 | 238 | 264 | 255 | 250 | 255 | 255 |  |  | 16
| 1 | 4 | 252 | 263.8554 | 260.6 | 249 | 253 | 247 | 268 | 256 | 238 | 248 | 258 | 260 | 253 | 240 | 231.1 | 270.4 | 232 | 234 | 281 | 260.3 | 236 | 287 | 240 | 261 | 253 | 250 | 253 | 253.8 |  |  | 252.1642
| 1 | 5 | 251 | 261.6413 | 260.7 | 253 | 251 | 252 | 271 | 274 | 243 | 253 | 259 | 238.9 | 259 | 244 | 234.4 | 269.5 | 232 | 227 | 272 | 268.1 | 239 | 275 | 248 | 256 | 241 | 252 | 249 | 256.7 |  |  | 66
| 1 | 6 | 251 | 265.2567 | 260.3 | 247 | 253 | 248 | 278 | 270 | 249 | 251 | 259 | 263.1 | 256 | 250 | 229.6 | 273.5 | 232 | 238 | 272 | 262.3 | 239 | 279 | 244 | 253 | 232 | 262 | 256 | 257.5
| Mean |  | 250.8333 | 264.3572 | 261.4833 | 247.5 | 252.6667 | 248.6667 | 273.1667 | 262.5 | 247.8333 | 253.6667 | 247.8333 | 251.4833 | 257.5 | 244.6667 | 237.4333 | 271.6 | 234.3333 | 232.6667 | 280.1667 | 264.4 | 241.3333 | 283 | 241 | 257.8333 | 247.1667 | 251.5 | 254.3333 | 254.5667
| Median |  | 251 | 264.2996 | 260.65 | 248 | 253 | 248.5 | 272.5 | 260 | 248.5 | 252 | 258.5 | 251.05 | 257 | 244 | 237.4 | 269.95 | 232.5 | 233 | 277.5 | 263.6 | 239 | 281 | 242 | 257 | 250 | 250 | 254 | 255.85
| Std Dev. |  | 5.5648 | 1.8043 | 3.3796 | 4.7645 | 2.6583 | 2.582 | 4.2151 | 7.7395 | 7.3598 | 7.4476 | 26.8732 | 13.7278 | 3.8341 | 3.4448 | 6.8529 | 3.5009 | 3.3862 | 3.6697 | 9.2177 | 3.3166 | 6.1536 | 8.0994 | 5.6214 | 4.0702 | 9.2177 | 5.4314 | 5.5737 | 3.7233
| Rel.Std.Dev. |  | 2.219% | 0.6825% | 1.292% | 1.925% | 1.052% | 1.038% | 1.543% | 2.948% | 2.97% | 2.936% | 10.84% | 5.459% | 1.489% | 1.408% | 2.886% | 1.289% | 1.445% | 1.577% | 3.29% | 1.254% | 2.55% | 2.862% | 2.333% | 1.579% | 3.729% | 2.16% | 2.192% | 1.463%
| PDM3 |  | -0.5278% | 4.835% | 3.696% | -1.85% | 0.1992% | -1.387% | 8.329% | 4.099% | -1.717% | 0.5958% | -1.717% | -0.27% | 2.116% | -2.973% | -5.842% | 7.708% | -7.071% | -7.732% | 11.1% | 4.852% | -4.295% | 12.23% | -4.427% | 2.248% | -1.982% | -0.2634% | 0.8602% | 0.9527%
| Z-Score (Absolute) |  | 0.12 | 1.17 | 0.9 | 0.44 | 0.06 | 0.33 | 2.01 | 0.99 | 0.4 | 0.12 | 0.64 | 0.06 | 0.52 | 0.71 | 1.4 | 1.86 | 1.8 | 1.85 | 2.68 | 1.18 | 1.25 | 2.95 | 1.06 | 0.55 | 0.47 | 0.26 | 0.22 | 0.37
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion ICP
| Analytical results for Ag in OREAS 905 (Indicative Value 0.685 ppm) | [24 empty] | Indicative
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | [17 empty] | 1
| No | No | 05 | 10 | 15 | 19 | 26 | 32 | [17 empty] | ppm
|  |  | PF*MS | PF*OES/MS | PF*OES/MS | PF*OES | ABL*MS | PF*OES/MS | [17 empty] | 3
3
| 1 | 1 | < 10 | 5 | < 1 | 0.63 | 0.9 | < 1 | [17 empty] | 1
| 1 | 2 | < 10 | 4 | < 1 | 0.64 | 0.9 | < 1 | [17 empty] | 6
| 1 | 3 | < 10 | 3 | < 1 | 0.61 | 0.4 | < 1 | [17 empty] | 16
| 1 | 4 | < 10 | 7 | < 1 | 0.6 | 0.7 | < 1 | [17 empty] | 0.685
| 1 | 5 | < 10 | 3 | < 1 | 0.6 | 0.8 | < 1 | [17 empty] | 12
| 1 | 6 | < 10 | 3 | < 1 | 0.64 | 0.8 | < 1
| Mean |  |  | 4.1667 |  | 0.62 | 0.75
| Median |  |  | 3.5 |  | 0.62 | 0.8
| Std Dev. |  |  | 1.6021 |  | 0.018974 | 0.18708
| Rel.Std.Dev. |  |  | 38.45% |  | 3.06% | 24.94%
| PDM3 |  |  | 508.3% |  | -9.489% | 9.489%
| Z-Score (Absolute) |  | NA | 17.72 | NA | 0.67 | 0 | NA
| Analytical results for Al in OREAS 905 (Certified Value 7.59 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 26 | 30 | 31 | 32 | 33 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | PF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*AAS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | 7.648 | 7.6966 | 7.63 | 7.19 | 7.992 | 7.7 | 7.56 | 7.23 | 7.67 | 7.18 | 7.53 | 6.12 | 7.37 | 7.43 | 7.53 | 7.545 | 7.7721 | 7.45 | 7.51 | 6.91 | 7.542 | 7.35 | 7.96 | 1
| 1 | 2 | 7.7 | 7.7396 | 7.74 | 7.37 | 7.859 | 7.886 | 7.7 | 7.79 | 7.62 | 7.17 | 7.58 | 7.17 | 7.25 | 7.45 | 7.44 | 7.608 | 7.7207 | 6.81 | 7.41 | 6.94 | 7.621 | 7.19 | 7.98 | 0
| 1 | 3 | 7.542 | 7.7929 | 7.83 | 7.66 | 7.833 | 7.7 | 7.54 | 7.64 | 7.58 | 7.01 | 7.56 | 6.01 | 7.45 | 7.34 | 7.61 | 7.506 | 7.7194 | 7.07 | 7.51 | 7.1 | 7.568 | 7.25 | 7.61 | 16
| 1 | 4 | 7.568 | 7.7184 | 7.58 | 7.53 | 7.859 | 7.78 | 7.53 | 7.49 | 7.64 | 6.94 | 7.6 | 6.87 | 7.5 | 7.39 | 7.76 | 7.3 | 7.6358 | 7.34 | 7.62 | 6.87 | 7.727 | 7.3 | 7.35 | 7.5948
| 1 | 5 | 7.595 | 7.5668 | 7.83 | 7.46 | 7.859 | 7.965 | 7.64 | 7.42 | 7.64 | 6.96 | 7.61 | 6.54 | 7.46 | 7.35 | 7.63 | 7.375 | 7.6546 | 7.14 | 7.57 | 6.75 | 7.648 | 7.41 | 7.63 | 68
| 1 | 6 | 7.595 | 7.6051 | 7.53 | 7.57 | 7.939 | 7.833 | 7.49 | 7.36 | 7.76 | 6.93 | 7.64 | 7.04 | 7.46 | 7.5 | 7.78 | 7.664 | 7.7012 | 6.97 | 7.62 | 6.74 | 7.727 | 7.39 | 7.7
| Mean |  | 7.608 | 7.6866 | 7.69 | 7.4633 | 7.8902 | 7.8107 | 7.5767 | 7.4883 | 7.6517 | 7.0317 | 7.5867 | 6.625 | 7.415 | 7.41 | 7.625 | 7.4997 | 7.7006 | 7.13 | 7.54 | 6.885 | 7.6388 | 7.315 | 7.705
| Median |  | 7.595 | 7.7075 | 7.685 | 7.495 | 7.859 | 7.8065 | 7.55 | 7.455 | 7.64 | 6.985 | 7.59 | 6.705 | 7.455 | 7.41 | 7.62 | 7.5255 | 7.7103 | 7.105 | 7.54 | 6.89 | 7.6345 | 7.325 | 7.665
| Std Dev. |  | 0.057198 | 0.085092 | 0.12884 | 0.16609 | 0.061542 | 0.10529 | 0.078145 | 0.20094 | 0.0608 | 0.11444 | 0.038816 | 0.48368 | 0.091378 | 0.061644 | 0.13096 | 0.13874 | 0.049362 | 0.23588 | 0.080498 | 0.13368 | 0.077901 | 0.084794 | 0.23721
| Rel.Std.Dev. |  | 0.7518% | 1.107% | 1.675% | 2.225% | 0.78% | 1.348% | 1.031% | 2.683% | 0.7946% | 1.628% | 0.5116% | 7.301% | 1.232% | 0.8319% | 1.717% | 1.85% | 0.641% | 3.308% | 1.068% | 1.942% | 1.02% | 1.159% | 3.079%
| PDM3 |  | 0.174% | 1.209% | 1.254% | -1.731% | 3.889% | 2.843% | -0.2385% | -1.402% | 0.749% | -7.415% | -0.1069% | -12.77% | -2.367% | -2.433% | 0.3979% | -1.252% | 1.394% | -6.12% | -0.7213% | -9.346% | 0.58% | -3.684% | 1.451%
| Z-Score (Absolute) |  | 0.19 | 0.65 | 0.67 | 0.67 | 1.86 | 1.39 | 0 | 0.53 | 0.45 | 3.24 | 0.06 | 5.66 | 0.96 | 0.99 | 0.29 | 0.46 | 0.74 | 2.66 | 0.22 | 4.12 | 0.37 | 1.56 | 0.76
| Analytical results for As in OREAS 905 (Certified Value 36.2 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 12 | 14 | 15 | 17 | 19 | 26 | 30 | 31 | 32 | 33 | [5 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | ABL*MS | PF*MS | PF*OES | PF*OES/MS | PF*OES | [5 empty] | 1
1
| 1 | 1 | < 100 | 36.0731 | < 30 | 30 | < 100 | < 100 | 18 | < 100 | 36 | 48 | 33 | 37 | 34.2 | 36.2 | 40 | 100 | 35.5 | < 70 | [5 empty] | 1
| 1 | 2 | < 100 | 36.2158 | < 30 | 30 | < 100 | < 100 | 17 | < 100 | 38 | 69 | 32 | 36 | 36.8 | 34.6 | 50 | 100 | 34.9 | < 70 | [5 empty] | 0
| 1 | 3 | < 100 | 37.408 | < 30 | < 30 | 100 | < 100 | 19 | < 100 | 36 | 90 | 33 | 40 | 35.5 | 32.4 | 30 | < 100 | 34.1 | < 70 | [5 empty] | 16
| 1 | 4 | < 100 | 34.2963 | < 30 | < 30 | < 100 | < 100 | 18 | < 100 | 52 | < 30 | 34 | 50 | 34.6 | 34.6 | 50 | 100 | 35.2 | < 70 | [5 empty] | 36.1983
| 1 | 5 | < 100 | 34.9967 | < 30 | 40 | < 100 | < 100 | 17 | < 100 | 37 | 43 | 34 | 37 | 34.5 | 37.8 | 30 | 200 | 33.7 | < 70 | [5 empty] | 69
| 1 | 6 | < 100 | 34.0201 | < 30 | < 30 | 100 | < 100 | 18 | < 100 | 39 | 92 | 33 | 48 | 35.9 | 35 | 20 | < 100 | 36 | < 70
| Mean |  | 03: 35.5016; 05: 33.3333; 06: 100; 09: 17.8333; 11: 39.6667; 12: 68.4; 13: 33.1667; 14: 41.3333; 15: 35.25; 17: 35.1; 18: 36.6667; 19: 125; 20: 34.9
| Median |  | 03: 35.5349; 05: 30; 06: 100; 09: 18; 11: 37.5; 12: 69; 13: 33; 14: 38.5; 15: 35.05; 17: 34.8; 18: 35; 19: 100; 20: 35.05
| Std Dev. |  | 03: 1.2941; 05: 5.7735; 06: 0; 09: 0.75277; 11: 6.1536; 12: 22.832; 13: 0.75277; 14: 6.121; 15: 0.9975; 17: 1.8055; 18: 12.1106; 19: 50; 20: 0.86487
| Rel.Std.Dev. |  | 03: 3.645%; 05: 17.32%; 06: 0%; 09: 4.221%; 11: 15.51%; 12: 33.38%; 13: 2.27%; 14: 14.81%; 15: 2.83%; 17: 5.144%; 18: 33.03%; 19: 40%; 20: 2.478%
| PDM3 |  | 03: -1.925%; 05: -7.915%; 06: 176.3%; 09: -50.73%; 11: 9.581%; 12: 88.96%; 13: -8.375%; 14: 14.19%; 15: -2.62%; 17: -3.034%; 18: 1.294%; 19: 245.3%; 20: -3.587%
| Z-Score (Absolute) |  | NA | 0 | NA | 0.63 | 18.63 | NA | 5.1 | NA | 0.49 | 9.5 | 0.67 | 1.68 | 0.07 | 0.12 | 1.3 | 25.85 | 0.17 | NA
| Analytical results for B in OREAS 905 (Indicative Value 19.2 ppm) | [24 empty] | Indicative
B | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 03; 03: 12; 04: 22; 05: 26 | ppm
|  |  | 01: PF*OES; 03: PF*MS; 04: PF*OES/MS; 05: PF*OES | 1
1
| 1 | 1 | 01: < 50; 03: 30; 04: 411; 05: < 20 | 1
| 1 | 2 | 01: < 50; 03: 30; 04: 422; 05: 20 | 3
| 1 | 3 | 01: < 50; 03: 20; 04: 420; 05: 20 | 16
| 1 | 4 | 01: < 50; 03: 20; 04: 405; 05: < 20 | 19.1667
| 1 | 5 | 01: < 50; 03: 10; 04: 394; 05: 20 | 13
| 1 | 6 | 01: < 50; 03: 20; 04: 359; 05: 20
| Mean |  | 03: 21.6667; 04: 401.8333; 05: 20
| Median |  | 03: 20; 04: 408; 05: 20
| Std Dev. |  | 03: 7.5277; 04: 23.3531; 05: 0
| Rel.Std.Dev. |  | 03: 34.74%; 04: 5.812%; 05: 0%
| PDM3 |  | 03: 13.04%; 04: 1997%; 05: 4.348%
| Z-Score (Absolute) |  | 01: NA; 03: 0; 04: 51.27; 05: 0.67
| Analytical results for Ba in OREAS 905 (Certified Value 2795 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 26 | 30 | 32 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | BF*MS | PF*OES/MS | BF*OES | PF*MS | BF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 2844.4784 | 2738 | 2850 | 3050 | 2770 | 2950 | 2723 | 2297 | 2810 | 2674 | 2771 | 2759.495 | > 2000 | 2960 | 2780 | 2770 | [7 empty] | 1
| 1 | 2 | 2830.1562 | 2748 | 2830 | 2950 | 2774 | 2880 | 2750 | 2666 | 2900 | 2715 | 2741 | 2682.624 | > 2000 | 3010 | 2810 | 2730 | [7 empty] | 0
| 1 | 3 | 2857.8053 | 2716 | 2630 | 3070 | 2725 | 2860 | 2742 | 2259 | 2850 | 2688 | 2673 | 2786.225 | > 2000 | 2990 | 2820 | 2720 | [7 empty] | 16
| 1 | 4 | 2887.5212 | 2719 | 2940 | 2870 | 2750 | 2870 | 2733 | 2589 | 2880 | 2682 | 2655 | 2720.259 | > 2000 | 2840 | 2800 | 2750 | [7 empty] | 2795.3371
| 1 | 5 | 2831.195 | 2712 | 2950 | 2970 | 2728 | 2630 | 2713 | 2424 | 2860 | 2664 | 2668 | 2683.747 | > 2000 | 3020 | 2740 | 2730 | [7 empty] | 70
| 1 | 6 | 2825.8749 | 2694 | 2710 | 2780 | 2765 | 2960 | 2735 | 2718 | 2860 | 2669 | 2671 | 2638.338 | > 2000 | 2950 | 2680 | 2760
| Mean |  | 2846.1718 | 2721.1667 | 2818.3333 | 2948.3333 | 2752 | 2858.3333 | 2732.6667 | 2492.1667 | 2860 | 2682 | 2696.5 | 2711.7813 |  | 2961.6667 | 2771.6667 | 2743.3333
| Median |  | 2837.8367 | 2717.5 | 2840 | 2960 | 2757.5 | 2875 | 2734 | 2506.5 | 2860 | 2678 | 2672 | 2702.003 |  | 2975 | 2790 | 2740
| Std Dev. |  | 23.4126 | 19.2709 | 126.8726 | 109.6206 | 21.3822 | 119.5687 | 13.2162 | 193.7518 | 30.3315 | 18.3412 | 47.471 | 54.6266 |  | 65.549 | 53.0723 | 19.6638
| Rel.Std.Dev. |  | 0.8226% | 0.7082% | 4.502% | 3.718% | 0.777% | 4.183% | 0.4836% | 7.774% | 1.061% | 0.6839% | 1.76% | 2.014% |  | 2.213% | 1.915% | 0.7168%
| PDM3 |  | 1.819% | -2.653% | 0.8227% | 5.473% | -1.55% | 2.254% | -2.242% | -10.85% | 2.313% | -4.055% | -3.536% | -2.989% |  | 5.95% | -0.8468% | -1.86%
| Z-Score (Absolute) |  | 0.91 | 0.3 | 0.64 | 1.89 | 0 | 1.46 | 0.19 | 2.5 | 1.04 | 0.67 | 0.68 | 0.39 | NA | 2.02 | 0.19 | 0.08
| Analytical results for Be in OREAS 905 (Certified Value 3.08 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 26 | 30 | 32 | [6 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [6 empty] | 2
2
| 1 | 1 | 3.442 | < 5 | 2.4 | 3.5 | < 5 | 2.8 | < 4 | 3 | < 5 | < 5 | 3 | 3.1 | 3.046 | 2.5 | 2.8 | < 5 | < 5 | [6 empty] | 1
| 1 | 2 | 3.3491 | < 5 | 2.4 | 4 | < 5 | 2.8 | < 4 | 2.9 | < 5 | < 5 | 4 | 3.4 | 3.22 | 2.2 | 4 | < 5 | < 5 | [6 empty] | 0
| 1 | 3 | 3.357 | < 5 | 2.3 | 3.5 | < 5 | 2.8 | < 4 | 2.9 | < 5 | < 5 | 3 | 3.3 | 3.688 | 2.5 | 2.8 | < 5 | < 5 | [6 empty] | 16
| 1 | 4 | 2.9963 | < 5 | 2.2 | 4 | < 5 | 2.9 | < 4 | 2.7 | < 5 | < 5 | 3 | 3.1 | 3.533 | 2.5 | 4 | < 5 | < 5 | [6 empty] | 3.0828
| 1 | 5 | 2.9438 | < 5 | 2.3 | 4 | < 5 | 2.8 | < 4 | 3 | < 5 | < 5 | 3 | 3 | 3.023 | 2.4 | 4 | < 5 | < 5 | [6 empty] | 71
| 1 | 6 | 2.9371 | < 5 | 2.6 | 4.5 | < 5 | 2.8 | < 4 | 3 | 5 | < 5 | 3 | 3 | 2.891 | 2.5 | 3.8 | < 5 | < 5
| Mean |  | 01: 3.1709; 04: 2.3667; 05: 3.9167; 07: 2.8167; 10: 2.9167; 11: 5; 13: 3.1667; 14: 3.15; 15: 3.2335; 17: 2.4333; 18: 3.5667
| Median |  | 01: 3.1727; 04: 2.35; 05: 4; 07: 2.8; 10: 2.95; 11: 5; 13: 3; 14: 3.1; 15: 3.133; 17: 2.5; 18: 3.9
| Std Dev. |  | 01: 0.23519; 04: 0.13663; 05: 0.37639; 07: 0.040825; 10: 0.1169; 13: 0.40825; 14: 0.16432; 15: 0.31408; 17: 0.12111; 18: 0.59889
| Rel.Std.Dev. |  | 01: 7.417%; 04: 5.773%; 05: 9.61%; 07: 1.449%; 10: 4.008%; 13: 12.89%; 14: 5.216%; 15: 9.713%; 17: 4.977%; 18: 16.79%
| PDM3 |  | 01: 2.858%; 04: -23.23%; 05: 27.05%; 07: -8.632%; 10: -5.388%; 11: 62.19%; 13: 2.721%; 14: 2.181%; 15: 4.889%; 17: -21.07%; 18: 15.7%
| Z-Score (Absolute) |  | 0.01 | NA | 1.54 | 1.44 | NA | 0.67 | NA | 0.4 | 3.53 | NA | 0 | 0.03 | 0.13 | 1.32 | 0.77 | NA | NA
| Analytical results for Bi in OREAS 905 (Certified Value 6.2 ppm) | [24 empty] | Certified
Bi | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 15; 09: 17; 10: 19; 11: 20; 12: 26; 13: 30; 14: 32 | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 09: PF*MS; 10: PF*OES; 11: PF*MS; 12: ABL*MS; 13: PF*MS; 14: PF*OES/MS | 2
3
| 1 | 1 | 01: 6.2054; 03: 6.3; 04: 6; 05: 4.6; 06: 5; 07: 6.5; 09: 5.9; 10: 7.3; 11: 6.185; 12: 6.08; 13: 2.9; 14: 5.8 | 1
| 1 | 2 | 01: 6.0289; 03: 6.2; 04: 6; 05: 5.9; 06: 5; 07: 6; 09: 6; 10: 7.1; 11: 6.128; 12: 6.68; 13: 3.6; 14: 5.87 | 0
| 1 | 3 | 01: 5.8889; 03: 6.2; 04: 6; 05: 8.4; 06: 4; 07: 6.4; 09: 6; 10: 7; 11: 5.997; 12: 6.18; 13: 3.2; 14: 5.88 | 16
| 1 | 4 | 01: 6.7776; 03: 6; 04: 6; 05: 39.2; 06: 5; 07: 6.7; 09: 5.8; 10: 6.3; 11: 6.248; 12: 5.66; 13: 2.9; 14: 5.88 | 6.1971
| 1 | 5 | 01: 6.4008; 03: 6; 04: 6; 05: 9.7; 06: 4; 07: 6.3; 09: 5.8; 10: 7.4; 11: 6.07; 12: 6.34; 13: 2.5; 14: 5.76 | 72
| 1 | 6 | 01: 6.444; 03: 6.1; 04: 6; 05: 9.8; 06: 5; 07: 6.5; 09: 5.9; 10: 6.8; 11: 6.462; 12: 5.88; 13: 2.7; 14: 5.8
| Mean |  | 01: 6.2909; 03: 6.1333; 04: 6; 05: 12.9333; 06: 4.6667; 07: 6.4; 09: 5.9; 10: 6.9833; 11: 6.1817; 12: 6.1367; 13: 2.9667; 14: 5.8317
| Median |  | 01: 6.3031; 03: 6.15; 04: 6; 05: 9.05; 06: 5; 07: 6.45; 09: 5.9; 10: 7.05; 11: 6.1565; 12: 6.13; 13: 2.9; 14: 5.835
| Std Dev. |  | 01: 0.31936; 03: 0.12111; 04: 0; 05: 13.0359; 06: 0.5164; 07: 0.23664; 09: 0.089443; 10: 0.39707; 11: 0.16276; 12: 0.3563; 13: 0.38816; 14: 0.051543
| Rel.Std.Dev. |  | 01: 5.076%; 03: 1.975%; 04: 0%; 05: 100.8%; 06: 11.07%; 07: 3.698%; 09: 1.516%; 10: 5.686%; 11: 2.633%; 12: 5.806%; 13: 13.08%; 14: 0.8838%
| PDM3 |  | 01: 1.514%; 03: -1.03%; 04: -3.181%; 05: 108.7%; 06: -24.7%; 07: 3.273%; 09: -4.795%; 10: 12.69%; 11: -0.2497%; 12: -0.9758%; 13: -52.13%; 14: -5.897%
| Z-Score (Absolute) |  | 01: 0.42; 03: 0; 04: 0.36; 05: 4.17; 06: 3.96; 07: 0.71; 09: 0.63; 10: 2.06; 11: 0.13; 12: 0; 13: 8.55; 14: 0.82
| Analytical results for Ca in OREAS 905 (Certified Value 0.608 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 3
3
| 1 | 1 | 0.622 | 0.63034 | 0.59 | 0.6 | 0.572 | 0.6 | 0.6 | 0.63 | 0.586 | 0.63 | 0.59 | 0.51 | 0.6 | 0.6 | 0.79 | 0.6103 | 0.849 | 0.62 | 0.622 | 0.589 | 0.593 |  |  | 1
| 1 | 2 | 0.629 | 0.6745 | 0.58 | 0.6 | 0.507 | 0.593 | 0.62 | 0.62 | 0.586 | 0.63 | 0.59 | 0.6 | 0.6 | 0.6 | 0.715 | 0.60675 | 0.801 | 0.62 | 0.622 | 0.578 | 0.609 |  |  | 0
| 1 | 3 | 0.679 | 0.62341 | 0.61 | 0.6 | 0.536 | 0.6 | 0.59 | 0.64 | 0.586 | 0.63 | 0.59 | 0.49 | 0.6 | 0.6 | 0.72 | 0.61922 | 0.925 | 0.63 | 0.65 | 0.578 | 0.569 |  |  | 16
| 1 | 4 | 0.65 | 0.66761 | 0.61 | 0.6 | 0.536 | 0.622 | 0.61 | 0.7 | 0.586 | 0.65 | 0.62 | 0.58 | 0.6 | 0.6 | 0.698 | 0.60591 | 0.871 | 0.61 | 0.658 | 0.602 | 0.527 |  |  | 0.60822
| 1 | 5 | 0.672 | 0.62756 | 0.62 | 0.6 | 0.557 | 0.658 | 0.6 | 0.68 | 0.586 | 0.6 | 0.59 | 0.53 | 0.6 | 0.6 | 0.73 | 0.60196 | 0.844 | 0.59 | 0.679 | 0.609 | 0.552 |  |  | 73
| 1 | 6 | 0.643 | 0.62301 | 0.59 | 0.6 | 0.536 | 0.6 | 0.58 | 0.62 | 0.586 | 0.65 | 0.59 | 0.6 | 0.6 | 0.6 | 0.77 | 0.60917 | 0.825 | 0.59 | 0.6 | 0.599 | 0.577
| Mean |  | 0.64917 | 0.64107 | 0.6 | 0.6 | 0.54067 | 0.61217 | 0.6 | 0.64833 | 0.586 | 0.63167 | 0.595 | 0.55167 | 0.6 | 0.6 | 0.73717 | 0.60889 | 0.8525 | 0.61 | 0.6385 | 0.5925 | 0.57117
| Median |  | 0.6465 | 0.62895 | 0.6 | 0.6 | 0.536 | 0.6 | 0.6 | 0.635 | 0.586 | 0.63 | 0.59 | 0.555 | 0.6 | 0.6 | 0.725 | 0.60796 | 0.8465 | 0.615 | 0.636 | 0.594 | 0.573
| Std Dev. |  | 0.022781 | 0.023486 | 0.015492 | 0 | 0.022124 | 0.024531 | 0.014142 | 0.033714 | 0 | 0.018348 | 0.012247 | 0.047924 | 0 | 0 | 0.035329 | 0.0058341 | 0.042642 | 0.016733 | 0.028912 | 0.012942 | 0.029206
| Rel.Std.Dev. |  | 3.509% | 3.663% | 2.582% | 0% | 4.092% | 4.007% | 2.357% | 5.2% | 0% | 2.905% | 2.058% | 8.687% | 0% | 0% | 4.793% | 0.9582% | 5.002% | 2.743% | 4.528% | 2.184% | 5.113%
| PDM3 |  | 6.732% | 5.401% | -1.352% | -1.352% | -11.11% | 0.6484% | -1.352% | 6.595% | -3.654% | 3.854% | -2.174% | -9.299% | -1.352% | -1.352% | 21.2% | 0.1088% | 40.16% | 0.2922% | 4.978% | -2.585% | -6.093%
| Z-Score (Absolute) |  | 2.38 | 1.99 | 0 | 0 | 2.86 | 0.15 | 0 | 2.34 | 0.67 | 1.53 | 0.48 | 1.09 | 0 | 0 | 6.63 | 0.43 | 12.2 | 0.48 | 1.86 | 0.36 | 1.39
| Analytical results for Cd in OREAS 905 (Indicative Value 0.4 ppm) | [24 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 10 | 12 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | [10 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [10 empty] | 2
2
| 1 | 1 | < 10 | < 10 | < 10 | 0.8 | < 2 | < 10 | 0.3 | < 10 | < 2 | 1 | 0.5 | < 10 | 0.33 | [10 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | 0.9 | < 2 | < 10 | 0.4 | < 10 | < 2 | < 1 | 0.4 | < 10 | 0.29 | [10 empty] | 8
| 1 | 3 | < 10 | < 10 | < 10 | 0.9 | < 2 | < 10 | 0.3 | < 10 | < 2 | < 1 | < 0.1 | < 10 | 0.31 | [10 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | 1 | < 2 | < 10 | 0.4 | < 10 | < 2 | 1 | < 0.1 | < 10 | 0.35 | [10 empty] | 0.4025
| 1 | 5 | < 10 | < 10 | < 10 | 1.1 | < 2 | < 10 | 0.4 | < 10 | < 2 | < 1 | 0.3 | < 10 | 0.36 | [10 empty] | 14
| 1 | 6 | < 10 | < 10 | < 10 | 0.9 | < 2 | < 10 | 0.3 | < 10 | < 2 | < 1 | 0.2 | < 10 | 0.42
| Mean |  | 05: 0.93333; 09: 0.35; 12: 1; 13: 0.35; 15: 0.34333
| Median |  | 05: 0.9; 09: 0.35; 12: 1; 13: 0.35; 15: 0.34
| Std Dev. |  | 05: 0.10328; 09: 0.054772; 12: 0; 13: 0.1291; 15: 0.045461
| Rel.Std.Dev. |  | 05: 11.07%; 09: 15.65%; 12: 0%; 13: 36.89%; 15: 13.24%
| PDM3 |  | 05: 131.9%; 09: -13.04%; 12: 148.4%; 13: -13.04%; 15: -14.7%
| Z-Score (Absolute) |  | NA | NA | NA | 3.93 | NA | NA | 0 | NA | NA | 2.14 | 0.67 | NA | 0.04
| Analytical results for Ce in OREAS 905 (Certified Value 96 ppm) | [24 empty] | Certified
Ce | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 11; 06: 12; 07: 13; 09: 15; 10: 20; 11: 22; 12: 26; 13: 30; 14: 32 |  |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: BF*OES; 06: PF*MS; 07: BF*MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES/MS; 12: ABL*MS; 13: PF*OES; 14: PF*OES/MS |  |  | 0
1
| 1 | 1 | 01: 94.9; 03: 91.5; 04: 99.7; 05: 96; 06: 91.5; 07: 92.5; 09: 101; 10: 94.415; 11: 105; 12: 96.1; 13: 90; 14: 99.2 |  |  | 1
| 1 | 2 | 01: 96.7; 03: 99; 04: 102; 05: 94; 06: 87.8; 07: 93.6; 09: 93.4; 10: 92.342; 11: 87; 12: 97.7; 13: 100; 14: 98.3 |  |  | 0
| 1 | 3 | 01: 89.3; 03: 91.5; 04: 107; 05: 101; 06: 86.2; 07: 92.2; 09: 97; 10: 97.168; 11: 92; 12: 102; 13: 100; 14: 97.9 |  |  | 16
| 1 | 4 | 01: 90.4; 03: 98.5; 04: 106; 05: 95; 06: 87.3; 07: 94.1; 09: 102; 10: 95.297; 11: 100; 12: 92.3; 13: 120; 14: 96.4 |  |  | 95.8806
| 1 | 5 | 01: 93.4; 03: 101; 04: 107; 05: 73; 06: 79.5; 07: 93.3; 09: 102; 10: 92.004; 11: 93; 12: 101; 13: 80; 14: 99.5 |  |  | 74
| 1 | 6 | 01: 88.3; 03: 106; 04: 109; 05: 96; 06: 93; 07: 93.4; 09: 101; 10: 90.878; 11: 98; 12: 97.8; 13: 90; 14: 98.5
| Mean |  | 01: 92.1667; 03: 97.9167; 04: 105.1167; 05: 92.5; 06: 87.55; 07: 93.1833; 09: 99.4; 10: 93.684; 11: 95.8333; 12: 97.8167; 13: 96.6667; 14: 98.3
| Median |  | 01: 91.9; 03: 98.75; 04: 106.5; 05: 95.5; 06: 87.55; 07: 93.35; 09: 101; 10: 93.3785; 11: 95.5; 12: 97.75; 13: 95; 14: 98.4
| Std Dev. |  | 01: 3.3417; 03: 5.634; 04: 3.5216; 05: 9.8539; 06: 4.734; 07: 0.70828; 09: 3.4756; 10: 2.3568; 11: 6.4317; 12: 3.4937; 13: 13.6626; 14: 1.1009
| Rel.Std.Dev. |  | 01: 3.626%; 03: 5.754%; 04: 3.35%; 05: 10.65%; 06: 5.407%; 07: 0.7601%; 09: 3.497%; 10: 2.516%; 11: 6.711%; 12: 3.572%; 13: 14.13%; 14: 1.12%
| PDM3 |  | 01: -3.874%; 03: 2.124%; 04: 9.633%; 05: -3.526%; 06: -8.689%; 07: -2.813%; 09: 3.671%; 10: -2.291%; 11: -0.04931%; 12: 2.019%; 13: 0.8198%; 14: 2.523%
| Z-Score (Absolute) |  | 01: 0.99; 03: 0.45; 04: 2.26; 05: 0.07; 06: 2.15; 07: 0.74; 09: 1.13; 10: 0.61; 11: 0.07; 12: 0.43; 13: 1.03; 14: 0.55
| Analytical results for Co in OREAS 905 (Certified Value 15.3 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  | 1
1
| 1 | 1 | 20 | 16.2089 | < 10 | 15 | 20 | < 20 | 20 | 11.9 | 14 | 17 | 15.7 | < 10 | 14.9 | 15 | 16.1 | 10 | 16.4 | 10 | < 50 | 15.6 | 20 |  |  | 1
| 1 | 2 | 20 | 16.2587 | < 10 | 15 | < 20 | < 20 | 20 | 11.9 | 15 | 16 | 15.5 | < 10 | 14.1 | 15 | 16.3 | 8 | 17.4 | < 10 | < 50 | 15.5 | 20 |  |  | 0
| 1 | 3 | < 20 | 16.0534 | < 10 | 15 | < 20 | 20 | 20 | 12.5 | 15 | 17.7 | 15 | < 10 | 14.8 | 15 | 15.9 | 9 | 14.4 | 10 | < 50 | 15.6 | 10 |  |  | 16
| 1 | 4 | < 20 | 15.967 | < 10 | 15 | < 20 | < 20 | 20 | 13 | 14 | 17.9 | 15.5 | < 10 | 15 | 15 | 15.7 | 10 | 15.3 | < 10 | < 50 | 15.3 | 20 |  |  | 15.2807
| 1 | 5 | 20 | 16.0386 | < 10 | 15 | < 20 | < 20 | 20 | 12.7 | 15 | 16.3 | 15.2 | < 10 | 15.1 | 15 | 15.6 | 9 | 15.7 | 20 | < 50 | 15.2 | 10 |  |  | 75
| 1 | 6 | < 20 | 16.1454 | < 10 | 15 | < 20 | < 20 | 20 | 13.2 | 14 | 13.3 | 15 | < 10 | 14.8 | 15 | 16.3 | 9 | 15.5 | 20 | < 50 | 15.6 | 20
| Mean |  | 20 | 16.112 |  | 15 | 20 | 20 | 20 | 12.5333 | 14.5 | 16.3667 | 15.3167 |  | 14.7833 | 15 | 15.9833 | 9.1667 | 15.7833 | 15 |  | 15.4667 | 16.6667
| Median |  | 20 | 16.0994 |  | 15 | 20 | 20 | 20 | 12.6 | 14.5 | 16.65 | 15.35 |  | 14.85 | 15 | 16 | 9 | 15.6 | 15 |  | 15.55 | 20
| Std Dev. |  | 0 | 0.11123 |  | 0 |  |  | 0 | 0.5465 | 0.54772 | 1.6777 | 0.29269 |  | 0.35449 | 0 | 0.29944 | 0.75277 | 1.0226 | 5.7735 |  | 0.17512 | 5.164
| Rel.Std.Dev. |  | 0% | 0.6903% |  | 0% |  |  | 0% | 4.36% | 3.777% | 10.25% | 1.911% |  | 2.398% | 0% | 1.873% | 8.212% | 6.479% | 38.49% |  | 1.132% | 30.98%
| PDM3 |  | 30.88% | 5.44% |  | -1.837% | 30.88% | 30.88% | 30.88% | -17.98% | -5.109% | 7.107% | 0.2356% |  | -3.255% | -1.837% | 4.598% | -40.01% | 3.29% | -1.837% |  | 1.217% | 9.07%
| Z-Score (Absolute) |  | 4.08 | 0.45 | NA | 0.58 | 4.08 | 4.08 | 4.08 | 2.88 | 1.05 | 0.69 | 0.29 | NA | 0.66 | 0.58 | 0.33 | 6.02 | 0.15 | 0.58 | NA | 0.15 | 0.97
| Analytical results for Cr in OREAS 905 (Indicative Value 28.2 ppm) | [24 empty] | Indicative
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 11 | 13 | 14 | 15 | 19 | 22 | 26 | 30 | 31 | 32 | 33 | [7 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | [7 empty] | 1
1
| 1 | 1 | < 50 | < 10 | < 20 | 40 | 30 | < 70 | 23 | 23 | 30 | 42.5 | 18 | < 50 | 30 | 70 | 19.9 | < 60 | [7 empty] | 1
| 1 | 2 | < 50 | < 10 | 20 | 40 | 40 | < 70 | 25 | 30 | 20 | 42.2 | 15 | < 50 | 30 | < 70 | 21.3 | < 60 | [7 empty] | 9
| 1 | 3 | < 50 | < 10 | 20 | 40 | 40 | < 70 | 26 | 31 | 20 | 34.3 | 17 | < 50 | 30 | < 70 | 18.9 | < 60 | [7 empty] | 16
| 1 | 4 | < 50 | < 10 | < 20 | 40 | 40 | < 70 | 21 | 42 | 20 | 35.6 | 34 | < 50 | 30 | < 70 | 22.3 | < 60 | [7 empty] | 28.2183
| 1 | 5 | < 50 | < 10 | < 20 | 40 | 30 | < 70 | 32 | 27 | 20 | 33.4 | 18 | < 50 | 30 | < 70 | 20.4 | < 60 | [7 empty] | 15
| 1 | 6 | < 50 | < 10 | < 20 | 40 | 30 | < 70 | 26 | 37 | 20 | 43 | 20 | < 50 | 40 | < 70 | 20.7 | < 60
| Mean |  | 04: 20; 05: 40; 06: 35; 09: 25.5; 10: 31.6667; 11: 21.6667; 12: 38.5; 13: 20.3333; 15: 31.6667; 17: 70; 19: 20.5833
| Median |  | 04: 20; 05: 40; 06: 35; 09: 25.5; 10: 30.5; 11: 20; 12: 38.9; 13: 18; 15: 30; 17: 70; 19: 20.55
| Std Dev. |  | 04: 0; 05: 0; 06: 5.4772; 09: 3.7283; 10: 6.8605; 11: 4.0825; 12: 4.5166; 13: 6.8896; 15: 4.0825; 19: 1.1669
| Rel.Std.Dev. |  | 04: 0%; 05: 0%; 06: 15.65%; 09: 14.62%; 10: 21.66%; 11: 18.84%; 12: 11.73%; 13: 33.88%; 15: 12.89%; 19: 5.669%
| PDM3 |  | 04: -29.12%; 05: 41.75%; 06: 24.03%; 09: -9.633%; 10: 12.22%; 11: -23.22%; 12: 36.44%; 13: -27.94%; 15: 12.22%; 17: 148.1%; 19: -27.06%
| Z-Score (Absolute) |  | NA | NA | 0.94 | 0.67 | 0.27 | NA | 0.5 | 0 | 0.81 | 0.55 | 1.14 | NA | 0 | 2.97 | 0.9 | NA
| Analytical results for Cs in OREAS 905 (Certified Value 7.1 ppm) | [24 empty] | Certified
Cs | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 09: 15; 10: 17; 11: 20; 12: 26; 13: 30; 14: 32 |  |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: BF*MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*MS; 12: ABL*MS; 13: PF*MS; 14: PF*OES/MS |  |  | 2
3
| 1 | 1 | 01: 6.1972; 03: 6.7; 04: 8; 05: 6.9; 06: 7.5; 07: 7.2; 09: 7.3; 10: 7; 11: 6.552; 12: 7.62; 13: 4.6; 14: 6.99 |  |  | 1
| 1 | 2 | 01: 6.4237; 03: 7.1; 04: 7; 05: 6.9; 06: 7.7; 07: 7.2; 09: 7.1; 10: 7.5; 11: 6.619; 12: 7.71; 13: 4.5; 14: 7.11 |  |  | 0
| 1 | 3 | 01: 6.1603; 03: 6.7; 04: 7; 05: 7.2; 06: 7.2; 07: 7; 09: 7.1; 10: 7.4; 11: 6.846; 12: 7.68; 13: 4.6; 14: 7.12 |  |  | 16
| 1 | 4 | 01: 6.1925; 03: 6.9; 04: 7; 05: 7.4; 06: 7.2; 07: 7; 09: 7.2; 10: 7.8; 11: 6.921; 12: 7.07; 13: 4.2; 14: 7.01 |  |  | 7.0989
| 1 | 5 | 01: 6.3928; 03: 7; 04: 7; 05: 7.3; 06: 6.4; 07: 7.2; 09: 7.1; 10: 6.8; 11: 7.017; 12: 7.9; 13: 4.3; 14: 6.96 |  |  | 76
| 1 | 6 | 01: 6.1546; 03: 6.6; 04: 8; 05: 7.6; 06: 10; 07: 7.1; 09: 7.2; 10: 7.7; 11: 7.062; 12: 7.73; 13: 4.5; 14: 7.03
| Mean |  | 01: 6.2535; 03: 6.8333; 04: 7.3333; 05: 7.2167; 06: 7.6667; 07: 7.1167; 09: 7.1667; 10: 7.3667; 11: 6.8362; 12: 7.6183; 13: 4.45; 14: 7.0367
| Median |  | 01: 6.1948; 03: 6.8; 04: 7; 05: 7.25; 06: 7.35; 07: 7.15; 09: 7.15; 10: 7.45; 11: 6.8835; 12: 7.695; 13: 4.5; 14: 7.02
| Std Dev. |  | 01: 0.12141; 03: 0.19664; 04: 0.5164; 05: 0.27869; 06: 1.2258; 07: 0.098319; 09: 0.08165; 10: 0.39328; 11: 0.20924; 12: 0.28449; 13: 0.16432; 14: 0.065013
| Rel.Std.Dev. |  | 01: 1.941%; 03: 2.878%; 04: 7.042%; 05: 3.862%; 06: 15.99%; 07: 1.382%; 09: 1.139%; 10: 5.339%; 11: 3.061%; 12: 3.734%; 13: 3.693%; 14: 0.9239%
| PDM3 |  | 01: -11.91%; 03: -3.741%; 04: 3.303%; 05: 1.659%; 06: 7.998%; 07: 0.2506%; 09: 0.9549%; 10: 3.772%; 11: -3.701%; 12: 7.317%; 13: -37.31%; 14: -0.8764%
| Z-Score (Absolute) |  | 01: 2.87; 03: 1; 04: 0.62; 05: 0.24; 06: 0.19; 07: 0.08; 09: 0.08; 10: 0.73; 11: 0.99; 12: 1.9; 13: 8.71; 14: 0.34
| Analytical results for Cu in OREAS 905 (Certified Value 1571 ppm) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 31 | 32 | 33 |  |  |  | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  | 0
0
| 1 | 1 | 1590 | 1558.3492 | 1565 | 1540 | 1600 | 1570 | 1650 | 1591 | 1610 | 1567 | 1302 | 1500 | 1539 | 1643 | 1606.55 | 1540 | 1560 | 1600 | 1510 | 1480 |  |  |  | 1
| 1 | 2 | 1580 | 1609.03 | 1549 | 1560 | 1610 | 1640 | 1630 | 1605 | 1650 | 1563 | 1524 | 1460 | 1543 | 1624 | 1454.75 | 1600 | 1590 | 1600 | 1530 | 1480 |  |  |  | 0
| 1 | 3 | 1560 | 1610.7783 | 1538 | 1520 | 1570 | 1600 | 1580 | 1612 | 1550 | 1706 | 1278 | 1500 | 1528 | 1635 | 1568.6 | 1580 | 1610 | 1600 | 1530 | 1530 |  |  |  | 16
| 1 | 4 | 1550 | 1580.6803 | 1556 | 1700 | 1580 | 1610 | 1420 | 1595 | 1600 | 1694 | 1445 | 1510 | 1542 | 1617 | 1593.9 | 1660 | 1600 | 1600 | 1530 | 1500 |  |  |  | 1571.4521
| 1 | 5 | 1580 | 1644.1083 | 1631 | 1590 | 1570 | 1630 | 1360 | 1596 | 1470 | 1569 | 1381 | 1510 | 1537 | 1617 | 1530.65 | 1640 | 1550 | 1600 | 1540 | 1620 |  |  |  | 77
| 1 | 6 | 1560 | 1563.1417 | 1599 | 1510 | 1590 | 1610 | 1430 | 1621 | 1650 | 1602 | 1488 | 1500 | 1557 | 1660 | 1518 | 1580 | 1530 | 1500 | 1540 | 1550
| Mean |  | 1570 | 1594.348 | 1573 | 1570 | 1586.6667 | 1610 | 1511.6667 | 1603.3333 | 1588.3333 | 1616.8333 | 1403 | 1496.6667 | 1541 | 1632.6667 | 1545.4083 | 1600 | 1573.3333 | 1583.3333 | 1530 | 1526.6667
| Median |  | 1570 | 1594.8552 | 1560.5 | 1550 | 1585 | 1610 | 1505 | 1600.5 | 1605 | 1585.5 | 1413 | 1500 | 1540.5 | 1629.5 | 1549.625 | 1590 | 1575 | 1600 | 1530 | 1515
| Std Dev. |  | 15.4919 | 32.9186 | 35.1966 | 69.857 | 16.3299 | 24.4949 | 123.1936 | 11.5355 | 68.8234 | 66.0225 | 99.94 | 18.619 | 9.4868 | 16.8839 | 56.2415 | 43.8178 | 31.4113 | 40.8248 | 10.9545 | 53.5413
| Rel.Std.Dev. |  | 0.9867% | 2.065% | 2.238% | 4.449% | 1.029% | 1.521% | 8.15% | 0.7195% | 4.333% | 4.083% | 7.123% | 1.244% | 0.6156% | 1.034% | 3.639% | 2.739% | 1.996% | 2.578% | 0.716% | 3.507%
| PDM3 |  | -0.0924% | 1.457% | 0.0985% | -0.0924% | 0.9682% | 2.453% | -3.804% | 2.029% | 1.074% | 2.888% | -10.72% | -4.759% | -1.938% | 3.895% | -1.657% | 1.817% | 0.1197% | 0.7561% | -2.638% | -2.85%
| Z-Score (Absolute) |  | 0.07 | 0.48 | 0 | 0.66 | 0.31 | 0.84 | 0.71 | 0.69 | 0.34 | 0.99 | 3.87 | 1.74 | 0.73 | 1.35 | 0.63 | 0.61 | 0 | 0.23 | 0.98 | 1.06
| Analytical results for Dy in OREAS 905 (Certified Value 3.78 ppm) | [24 empty] | Certified
Dy | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 |  |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 01: 3.7; 03: 3.5; 04: 4.56; 05: 3.5; 06: 3.6; 07: 3.79; 09: 3.908; 10: 3.88; 11: 4.02 |  |  | 1
| 1 | 2 | 01: 3.4; 03: 3.5; 04: 4.18; 05: 3.4; 06: 3.7; 07: 3.79; 09: 3.518; 10: 4.27; 11: 4.01 |  |  | 0
| 1 | 3 | 01: 3.5; 03: 3; 04: 4.24; 05: 3.4; 06: 3.6; 07: 3.82; 09: 3.588; 10: 4.62; 11: 3.89 |  |  | 16
| 1 | 4 | 01: 3.5; 03: 4; 04: 4.36; 05: 3.4; 06: 3.7; 07: 3.99; 09: 3.63; 10: 3.81; 11: 3.91 |  |  | 3.7842
| 1 | 5 | 01: 3.5; 03: 3.5; 04: 4.45; 05: 3.2; 06: 3.7; 07: 3.84; 09: 3.392; 10: 4.35; 11: 3.85 |  |  | 78
| 1 | 6 | 01: 3.6; 03: 4; 04: 4.21; 05: 3.6; 06: 3.6; 07: 3.9; 09: 3.67; 10: 3.94; 11: 3.86
| Mean |  | 01: 3.5333; 03: 3.5833; 04: 4.3333; 05: 3.4167; 06: 3.65; 07: 3.855; 09: 3.6177; 10: 4.145; 11: 3.9233
| Median |  | 01: 3.5; 03: 3.5; 04: 4.3; 05: 3.4; 06: 3.65; 07: 3.83; 09: 3.609; 10: 4.105; 11: 3.9
| Std Dev. |  | 01: 0.10328; 03: 0.37639; 04: 0.15042; 05: 0.13292; 06: 0.054772; 07: 0.077653; 09: 0.17258; 10: 0.31867; 11: 0.074207
| Rel.Std.Dev. |  | 01: 2.923%; 03: 10.5%; 04: 3.471%; 05: 3.89%; 06: 1.501%; 07: 2.014%; 09: 4.771%; 10: 7.688%; 11: 1.891%
| PDM3 |  | 01: -6.629%; 03: -5.308%; 04: 14.51%; 05: -9.712%; 06: -3.546%; 07: 1.871%; 09: -4.4%; 10: 9.535%; 11: 3.677%
| Z-Score (Absolute) |  | 01: 0.38; 03: 0.22; 04: 2.25; 05: 0.77; 06: 0; 07: 0.67; 09: 0.11; 10: 1.63; 11: 0.9
| Analytical results for Er in OREAS 905 (Certified Value 1.14 ppm) | [24 empty] | Certified
Er | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 |  |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS |  |  | 2
2
| 1 | 1 | 01: 1; 03: < 1; 04: 0.98; 05: 1; 06: 1.2; 07: 1.29; 09: 0.994; 10: 1.5; 11: 1.2 |  |  | 1
| 1 | 2 | 01: 1.2; 03: < 1; 04: 1; 05: 1; 06: 1.1; 07: 1.22; 09: 0.975; 10: 1.22; 11: 1.13 |  |  | 13
| 1 | 3 | 01: 1.2; 03: < 1; 04: 1.04; 05: 1.1; 06: 1.2; 07: 1.23; 09: 0.974; 10: 1.59; 11: 1.19 |  |  | 16
| 1 | 4 | 01: 1.3; 03: 2; 04: 1.09; 05: 1; 06: 1.2; 07: 1.3; 09: 1.088; 10: 1.19; 11: 1.12 |  |  | 1.1377
| 1 | 5 | 01: 1.1; 03: < 1; 04: 1.05; 05: 0.9; 06: 1.1; 07: 1.21; 09: 1.13; 10: 1.29; 11: 1.16 |  |  | 79
| 1 | 6 | 01: 1.2; 03: < 1; 04: 1.07; 05: 1; 06: 1.2; 07: 1.25; 09: 1.06; 10: 1.24; 11: 1.13
| Mean |  | 01: 1.1667; 03: 2; 04: 1.0383; 05: 1; 06: 1.1667; 07: 1.25; 09: 1.0368; 10: 1.3383; 11: 1.155
| Median |  | 01: 1.2; 03: 2; 04: 1.045; 05: 1; 06: 1.2; 07: 1.24; 09: 1.027; 10: 1.265; 11: 1.145
| Std Dev. |  | 01: 0.10328; 04: 0.041673; 05: 0.063246; 06: 0.05164; 07: 0.037417; 09: 0.065484; 10: 0.16582; 11: 0.033912
| Rel.Std.Dev. |  | 01: 8.853%; 04: 4.013%; 05: 6.325%; 06: 4.426%; 07: 2.993%; 09: 6.316%; 10: 12.39%; 11: 2.936%
| PDM3 |  | 01: 2.547%; 03: 75.8%; 04: -8.733%; 05: -12.1%; 06: 2.547%; 07: 9.872%; 09: -8.865%; 10: 17.64%; 11: 1.522%
| Z-Score (Absolute) |  | 01: 0; 03: 4.63; 04: 0.71; 05: 0.93; 06: 0; 07: 0.46; 09: 0.72; 10: 0.67; 11: 0.06
| Analytical results for Eu in OREAS 905 (Certified Value 1.46 ppm) | [24 empty] | Certified
Eu | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 |  |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS |  |  | 2
3
| 1 | 1 | 01: 1.4; 03: 1.5; 04: 2.5; 05: 1.4; 06: 1.4; 07: 1.46; 09: 1.416; 10: 1.35; 11: 1.44 |  |  | 1
| 1 | 2 | 01: 1.6; 03: 1.5; 04: 2.54; 05: 1.4; 06: 1.4; 07: 1.3; 09: 1.359; 10: 1.31; 11: 1.45 |  |  | 0
| 1 | 3 | 01: 1.7; 03: 1.5; 04: 2.68; 05: 1.4; 06: 1.4; 07: 1.46; 09: 1.558; 10: 1.65; 11: 1.4 |  |  | 16
| 1 | 4 | 01: 1.5; 03: 1.5; 04: 2.82; 05: 1.3; 06: 1.4; 07: 1.54; 09: 1.524; 10: 1.37; 11: 1.41 |  |  | 1.4579
| 1 | 5 | 01: 1.6; 03: 1.5; 04: 2.62; 05: 1.3; 06: 1.4; 07: 1.49; 09: 1.583; 10: 1.47; 11: 1.43 |  |  | 80
| 1 | 6 | 01: 1.6; 03: 1.5; 04: 2.77; 05: 1.4; 06: 1.4; 07: 1.5; 09: 1.477; 10: 1.46; 11: 1.38
| Mean |  | 01: 1.5667; 03: 1.5; 04: 2.655; 05: 1.3667; 06: 1.4; 07: 1.4583; 09: 1.4862; 10: 1.435; 11: 1.4183
| Median |  | 01: 1.6; 03: 1.5; 04: 2.65; 05: 1.4; 06: 1.4; 07: 1.475; 09: 1.5005; 10: 1.415; 11: 1.42
| Std Dev. |  | 01: 0.10328; 03: 0; 04: 0.12613; 05: 0.05164; 06: 2.4324e-16; 07: 0.083046; 09: 0.086194; 10: 0.1226; 11: 0.026394
| Rel.Std.Dev. |  | 01: 6.592%; 03: 0%; 04: 4.751%; 05: 3.779%; 06: 1.737e-14%; 07: 5.695%; 09: 5.8%; 10: 8.543%; 11: 1.861%
| PDM3 |  | 01: 7.464%; 03: 2.891%; 04: 82.12%; 05: -6.255%; 06: -3.968%; 07: 0.03287%; 09: 1.942%; 10: -1.568%; 11: -2.711%
| Z-Score (Absolute) |  | 01: 0.8; 03: 0.14; 04: 11.62; 05: 1.19; 06: 0.86; 07: 0.04; 09: 0; 10: 0.51; 11: 0.67
| Analytical results for Fe in OREAS 905 (Certified Value 4.23 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 26 | 30 | 31 | 32 | 33 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*AAS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | 2
3
| 1 | 1 | 4.302 | 4.1555 | 4.12 | 4.28 | 4.35 | 4.211 | 4.3 | 3.98 | 4.162 | 4.34 | 4.25 | 3.4 | 4.31 | 4.2 | 4.34 | 4.42 | 4.1688 | 4.209 | 4.15 | 3.81 | 4.358 | 4.23 | 3.97 | 1
| 1 | 2 | 4.26 | 4.2429 | 4.1 | 4.22 | 4.3 | 4.323 | 4.4 | 4 | 4.197 | 4.38 | 4.25 | 3.96 | 4.19 | 4.22 | 4.38 | 4.532 | 4.1959 | 3.7145 | 4.1 | 3.87 | 4.302 | 4.15 | 3.8 | 0
| 1 | 3 | 4.19 | 4.3093 | 4.23 | 4.16 | 4.26 | 4.26 | 4.23 | 4.14 | 4.253 | 4.25 | 4.21 | 3.36 | 4.22 | 4.13 | 4.25 | 4.438 | 4.1662 | 4.1975 | 4.11 | 3.99 | 4.302 | 4.14 | 4.09 | 16
| 1 | 4 | 4.218 | 4.1533 | 4.07 | 4.3 | 4.31 | 4.281 | 4.31 | 4.03 | 4.274 | 4.18 | 4.23 | 3.87 | 4.26 | 4.16 | 4.37 | 4.2 | 4.1093 | 4.232 | 4.16 | 3.82 | 4.372 | 4.17 | 3.83 | 4.2268
| 1 | 5 | 4.267 | 4.1571 | 4.1 | 4.43 | 4.24 | 4.323 | 4.51 | 3.83 | 4.197 | 4.21 | 4.23 | 3.61 | 4.28 | 4.16 | 4.29 | 4.25 | 4.1241 | 4.071 | 4.16 | 3.7 | 4.323 | 4.09 | 4.3 | 81
| 1 | 6 | 4.288 | 4.1598 | 4.02 | 4.15 | 4.31 | 4.267 | 4.31 | 3.89 | 4.239 | 4.2 | 4.25 | 4.19 | 4.21 | 4.22 | 4.25 | 4.335 | 4.193 | 4.071 | 4.12 | 3.75 | 4.372 | 4.16 | 4.14
| Mean |  | 4.2542 | 4.1963 | 4.1067 | 4.2567 | 4.295 | 4.2775 | 4.3433 | 3.9783 | 4.2203 | 4.26 | 4.2367 | 3.7317 | 4.245 | 4.1817 | 4.3133 | 4.3625 | 4.1595 | 4.0825 | 4.1333 | 3.8233 | 4.3382 | 4.1567 | 4.0217
| Median |  | 4.2635 | 4.1584 | 4.1 | 4.25 | 4.305 | 4.274 | 4.31 | 3.99 | 4.218 | 4.23 | 4.24 | 3.74 | 4.24 | 4.18 | 4.315 | 4.3775 | 4.1675 | 4.1342 | 4.135 | 3.815 | 4.3405 | 4.155 | 4.03
| Std Dev. |  | 0.04256 | 0.065314 | 0.069761 | 0.10443 | 0.03937 | 0.042406 | 0.097912 | 0.1087 | 0.041922 | 0.081731 | 0.01633 | 0.32993 | 0.045935 | 0.037103 | 0.058195 | 0.12454 | 0.03564 | 0.19353 | 0.026583 | 0.10073 | 0.033253 | 0.045461 | 0.19219
| Rel.Std.Dev. |  | 1% | 1.556% | 1.699% | 2.453% | 0.9166% | 0.9914% | 2.254% | 2.732% | 0.9933% | 1.919% | 0.3854% | 8.841% | 1.082% | 0.8873% | 1.349% | 2.855% | 0.8568% | 4.74% | 0.6431% | 2.635% | 0.7665% | 1.094% | 4.779%
| PDM3 |  | 0.6476% | -0.721% | -2.842% | 0.7068% | 1.614% | 1.2% | 2.757% | -5.878% | -0.1528% | 0.7856% | 0.2336% | -11.71% | 0.4308% | -1.068% | 2.047% | 3.211% | -1.591% | -3.414% | -2.211% | -9.545% | 2.635% | -1.659% | -4.853%
| Z-Score (Absolute) |  | 0.36 | 0.49 | 1.46 | 0.39 | 0.79 | 0.6 | 0.95 | 2.55 | 0 | 0.42 | 0.18 | 5.15 | 0.26 | 0.4 | 0.98 | 1.5 | 0.64 | 0.67 | 0.91 | 4.19 | 1.24 | 0.67 | 1
| Analytical results for Ga in OREAS 905 (Certified Value 25.8 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 05 | 09 | 10 | 12 | 13 | 15 | 17 | 19 | 20 | 22 | 26 | 30 | 32 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | [9 empty] | 1
2
| 1 | 1 | 26.3677 | 22 | 26 | 20 | 24.4 | 25.8 | 25 | 24 | 25.8 | 26.533 | 28 | 25.6 | < 2.5 | 27.3 | [9 empty] | 1
| 1 | 2 | 27.3849 | 21 | 25 | 20 | 24.5 | 25.5 | 24 | 26 | 26.6 | 24.666 | 25 | 26.1 | < 2.5 | 27.1 | [9 empty] | 0
| 1 | 3 | 27.2818 | 20 | 23 | 21 | 23.7 | 25.4 | 24 | 24 | 26.2 | 28.397 | 26 | 26.3 | < 2.5 | 27 | [9 empty] | 16
| 1 | 4 | 26.7387 | 22 | 27 | 21 | 23.6 | 25.4 | 26 | 25 | 26.2 | 28.188 | 28 | 25.1 | < 2.5 | 27 | [9 empty] | 25.8159
| 1 | 5 | 26.9511 | 20 | 26 | 21 | 21.7 | 25.4 | 25 | 23 | 26.2 | 27.189 | 27 | 27.7 | < 2.5 | 26.2 | [9 empty] | 82
| 1 | 6 | 28.0541 | 20 | 29 | 21 | 24.6 | 25.5 | 25 | 22 | 25.5 | 26.741 | 26 | 24.4 | < 2.5 | 26.1
| Mean |  | 01: 27.1297; 03: 20.8333; 04: 26; 05: 20.6667; 06: 23.75; 07: 25.5; 09: 24.8333; 10: 24; 11: 26.0833; 12: 26.9523; 13: 26.6667; 14: 25.8667; 17: 26.7833
| Median |  | 01: 27.1165; 03: 20.5; 04: 26; 05: 21; 06: 24.05; 07: 25.45; 09: 25; 10: 24; 11: 26.2; 12: 26.965; 13: 26.5; 14: 25.85; 17: 27
| Std Dev. |  | 01: 0.58449; 03: 0.98319; 04: 2; 05: 0.5164; 06: 1.0895; 07: 0.15492; 09: 0.75277; 10: 1.4142; 11: 0.38166; 12: 1.3504; 13: 1.2111; 14: 1.1325; 17: 0.50365
| Rel.Std.Dev. |  | 01: 2.154%; 03: 4.719%; 04: 7.692%; 05: 2.499%; 06: 4.587%; 07: 0.6075%; 09: 3.031%; 10: 5.893%; 11: 1.463%; 12: 5.01%; 13: 4.541%; 14: 4.378%; 17: 1.88%
| PDM3 |  | 01: 5.089%; 03: -19.3%; 04: 0.713%; 05: -19.95%; 06: -8.003%; 07: -1.224%; 09: -3.806%; 10: -7.034%; 11: 1.036%; 12: 4.402%; 13: 3.295%; 14: 0.1965%; 17: 3.747%
| Z-Score (Absolute) |  | 0.82 | 3.28 | 0.09 | 3.39 | 1.11 | 0.24 | 0.67 | 1.22 | 0.14 | 0.71 | 0.52 | 0 | NA | 0.6
| Analytical results for Gd in OREAS 905 (Certified Value 6.43 ppm) | [24 empty] | Certified
Gd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS | 2
2
| 1 | 1 | 01: 6.1; 03: 8; 04: 7.23; 05: 6.3; 06: 5.78; 07: 5.92; 09: 5.996; 10: 6; 11: 5.99 | 1
| 1 | 2 | 01: 5.8; 03: 8; 04: 7.31; 05: 6.2; 06: 5.95; 07: 5.81; 09: 5.472; 10: 6.64; 11: 5.9 | 0
| 1 | 3 | 01: 6; 03: 7; 04: 7.63; 05: 6.1; 06: 5.73; 07: 6.08; 09: 6.762; 10: 6.33; 11: 5.93 | 16
| 1 | 4 | 01: 6.1; 03: 8; 04: 7.87; 05: 6.1; 06: 5.88; 07: 6.32; 09: 5.457; 10: 5.9; 11: 6.04 | 6.4258
| 1 | 5 | 01: 6.1; 03: 8; 04: 7.64; 05: 5.4; 06: 5.85; 07: 6.06; 09: 6.329; 10: 6.39; 11: 6.18 | 83
| 1 | 6 | 01: 6.2; 03: 9; 04: 8.15; 05: 6.1; 06: 5.95; 07: 6.28; 09: 5.837; 10: 6.14; 11: 5.9
| Mean |  | 01: 6.05; 03: 8; 04: 7.6383; 05: 6.0333; 06: 5.8567; 07: 6.0783; 09: 5.9755; 10: 6.2333; 11: 5.99
| Median |  | 01: 6.1; 03: 8; 04: 7.635; 05: 6.1; 06: 5.865; 07: 6.07; 09: 5.9165; 10: 6.235; 11: 5.96
| Std Dev. |  | 01: 0.13784; 03: 0.63246; 04: 0.34354; 05: 0.32042; 06: 0.089368; 07: 0.19823; 09: 0.50705; 10: 0.27333; 11: 0.10807
| Rel.Std.Dev. |  | 01: 2.278%; 03: 7.906%; 04: 4.498%; 05: 5.311%; 06: 1.526%; 07: 3.261%; 09: 8.485%; 10: 4.385%; 11: 1.804%
| PDM3 |  | 01: -5.848%; 03: 24.5%; 04: 18.87%; 05: -6.108%; 06: -8.857%; 07: -5.407%; 09: -7.008%; 10: -2.995%; 11: -6.782%
| Z-Score (Absolute) |  | 01: 0; 03: 9.21; 04: 8.33; 05: 0.32; 06: 1.32; 07: 0.12; 09: 0.67; 10: 0.72; 11: 0.6
| Analytical results for Ge in OREAS 905 (Indicative Value 1.53 ppm) | [24 empty] | Indicative
Ge | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab | 1
| No | No | 01: 03; 03: 10; 04: 12; 05: 15; 06: 17; 07: 26; 09: 32 | ppm
|  |  | 01: PF*MS; 03: PF*OES/MS; 04: PF*MS; 05: PF*OES/MS; 06: PF*MS; 07: ABL*MS; 09: PF*OES/MS | 2
2
| 1 | 1 | 01: 1.5975; 03: 1; 04: 2.7; 05: 1; 06: 1; 07: 0.7; 09: 1.76 | 1
| 1 | 2 | 01: 1.7623; 03: 1; 04: 2.7; 05: 1; 06: 1; 07: 0.6; 09: 1.64 | 4
| 1 | 3 | 01: 2.0214; 03: 1; 04: 2.7; 05: 2; 06: 1; 07: 0.45; 09: 1.77 | 16
| 1 | 4 | 01: 2.136; 03: 1; 04: 2.5; 05: 2; 06: 2; 07: 1; 09: 1.66 | 1.5303
| 1 | 5 | 01: 1.9176; 03: 1; 04: 2.3; 05: 1; 06: 1; 07: 0.85; 09: 1.62 | 16
| 1 | 6 | 01: 1.4597; 03: 1; 04: 2.6; 05: 2; 06: 2; 07: 0.6; 09: 1.89
| Mean |  | 01: 1.8157; 03: 1; 04: 2.5833; 05: 1.5; 06: 1.3333; 07: 0.7; 09: 1.7233
| Median |  | 01: 1.84; 03: 1; 04: 2.65; 05: 1.5; 06: 1; 07: 0.65; 09: 1.71
| Std Dev. |  | 01: 0.25787; 03: 0; 04: 0.16021; 05: 0.54772; 06: 0.5164; 07: 0.19748; 09: 0.10289
| Rel.Std.Dev. |  | 01: 14.2%; 03: 0%; 04: 6.202%; 05: 36.51%; 06: 38.73%; 07: 28.21%; 09: 5.97%
| PDM3 |  | 01: 18.65%; 03: -34.66%; 04: 68.81%; 05: -1.983%; 06: -12.87%; 07: -54.26%; 09: 12.61%
| Z-Score (Absolute) |  | 01: 0.67; 03: 1.07; 04: 2.43; 05: 0; 06: 0.36; 07: 1.71; 09: 0.48
| Analytical results for Hf in OREAS 905 (Certified Value 7.26 ppm) | [24 empty] | Certified
Hf | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab | 1
| No | No | 01: 09; 03: 12; 04: 13; 05: 15; 06: 20; 07: 22; 09: 26; 10: 32 | ppm
|  |  | 01: BF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 09: ABL*MS; 10: PF*OES/MS | 2
2
| 1 | 1 | 01: 8; 03: < 10; 04: 6.9; 05: 8; 06: 6.681; 07: 33; 09: 7.96; 10: 7.56 | 1
| 1 | 2 | 01: 8; 03: < 10; 04: 7.1; 05: 6; 06: 6.577; 07: 51; 09: 7.93; 10: 7.47 | 0
| 1 | 3 | 01: 7; 03: < 10; 04: 7; 05: 8; 06: 6.23; 07: 7; 09: 8.2; 10: 7.3 | 16
| 1 | 4 | 01: 8; 03: < 10; 04: 7; 05: 8; 06: 6.375; 07: 7; 09: 7.84; 10: 7.36 | 7.2583
| 1 | 5 | 01: 8; 03: < 10; 04: 6.9; 05: 7; 06: 7.083; 07: 6; 09: 8.04; 10: 7.15 | 84
| 1 | 6 | 01: 9; 03: < 10; 04: 6.9; 05: 8; 06: 5.822; 07: 6; 09: 8; 10: 7.47
| Mean |  | 01: 8; 04: 6.9667; 05: 7.5; 06: 6.4613; 07: 18.3333; 09: 7.995; 10: 7.385
| Median |  | 01: 8; 04: 6.95; 05: 8; 06: 6.476; 07: 7; 09: 7.98; 10: 7.415
| Std Dev. |  | 01: 0.63246; 04: 0.08165; 05: 0.83666; 06: 0.42844; 07: 19.2007; 09: 0.12128; 10: 0.14707
| Rel.Std.Dev. |  | 01: 7.906%; 04: 1.172%; 05: 11.16%; 06: 6.631%; 07: 104.7%; 09: 1.517%; 10: 1.991%
| PDM3 |  | 01: 10.22%; 04: -4.018%; 05: 3.33%; 06: -10.98%; 07: 152.6%; 09: 10.15%; 10: 1.746%
| Z-Score (Absolute) |  | 01: 0.68; 03: NA; 04: 0.46; 05: 0.13; 06: 1.02; 07: 0.98; 09: 0.67; 10: 0
| Analytical results for Ho in OREAS 905 (Certified Value 0.53 ppm) | [24 empty] | Certified
Ho | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS | 2
2
| 1 | 1 | 01: 0.6; 03: < 1; 04: 0.41; 05: 0.5; 06: 0.5; 07: 0.61; 09: 0.522; 10: 0.61; 11: 0.578 | 1
| 1 | 2 | 01: 0.5; 03: < 1; 04: 0.38; 05: 0.5; 06: 0.5; 07: 0.56; 09: 0.566; 10: 0.65; 11: 0.579 | 0
| 1 | 3 | 01: 0.5; 03: < 1; 04: 0.42; 05: 0.5; 06: 0.5; 07: 0.57; 09: 0.451; 10: 0.66; 11: 0.568 | 16
| 1 | 4 | 01: 0.6; 03: < 1; 04: 0.45; 05: 0.5; 06: 0.5; 07: 0.66; 09: 0.562; 10: 0.58; 11: 0.593 | 0.53431
| 1 | 5 | 01: 0.6; 03: < 1; 04: 0.42; 05: 0.4; 06: 0.5; 07: 0.59; 09: 0.49; 10: 0.61; 11: 0.595 | 85
| 1 | 6 | 01: 0.5; 03: < 1; 04: 0.45; 05: 0.5; 06: 0.5; 07: 0.6; 09: 0.557; 10: 0.62; 11: 0.536
| Mean |  | 01: 0.55; 04: 0.42167; 05: 0.48333; 06: 0.5; 07: 0.59833; 09: 0.52467; 10: 0.62167; 11: 0.57483
| Median |  | 01: 0.55; 04: 0.42; 05: 0.5; 06: 0.5; 07: 0.595; 09: 0.5395; 10: 0.615; 11: 0.5785
| Std Dev. |  | 01: 0.054772; 04: 0.026394; 05: 0.040825; 06: 0; 07: 0.035449; 09: 0.04644; 10: 0.029269; 11: 0.021535
| Rel.Std.Dev. |  | 01: 9.959%; 04: 6.26%; 05: 8.447%; 06: 0%; 07: 5.925%; 09: 8.851%; 10: 4.708%; 11: 3.746%
| PDM3 |  | 01: 2.936%; 04: -21.08%; 05: -9.541%; 06: -6.422%; 07: 11.98%; 09: -1.805%; 10: 16.35%; 11: 7.584%
| Z-Score (Absolute) |  | 01: 0.19; 03: NA; 04: 1.7; 05: 0.8; 06: 0.55; 07: 0.9; 09: 0.19; 10: 1.24; 11: 0.55
| Analytical results for In in OREAS 905 (Certified Value 0.62 ppm) | [24 empty] | Certified
In | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 15; 09: 17; 10: 20; 11: 26; 12: 30; 13: 32 | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 09: PF*MS; 10: PF*MS; 11: ABL*MS; 12: PF*MS; 13: PF*OES/MS | 2
2
| 1 | 1 | 01: 0.59573; 03: 0.7; 04: 0.5; 05: 0.6; 06: < 0.2; 07: 0.7; 09: 0.6; 10: 0.75; 11: 0.6; 12: 0.3; 13: 0.69 | 1
| 1 | 2 | 01: 0.64888; 03: 0.7; 04: 0.5; 05: 0.6; 06: 0.4; 07: 0.6; 09: 0.6; 10: 0.611; 11: 0.6; 12: 0.3; 13: 0.65 | 0
| 1 | 3 | 01: 0.59158; 03: 0.6; 04: 0.5; 05: 0.6; 06: < 0.2; 07: 0.7; 09: 0.6; 10: 0.746; 11: 0.6; 12: 0.3; 13: 0.67 | 16
| 1 | 4 | 01: 0.42585; 03: 0.7; 04: 0.5; 05: 0.6; 06: < 0.2; 07: 0.7; 09: 0.7; 10: 0.613; 11: 0.55; 12: < 0.25; 13: 0.64 | 0.62198
| 1 | 5 | 01: 0.54245; 03: 0.6; 04: 0.5; 05: 0.6; 06: < 0.2; 07: 0.7; 09: 0.7; 10: 0.704; 11: 0.65; 12: < 0.25; 13: 0.64 | 86
| 1 | 6 | 01: 0.44358; 03: 0.7; 04: 1; 05: 0.6; 06: < 0.2; 07: 0.7; 09: 0.7; 10: 0.667; 11: 0.7; 12: 0.3; 13: 0.76
| Mean |  | 01: 0.54135; 03: 0.66667; 04: 0.58333; 05: 0.6; 06: 0.4; 07: 0.68333; 09: 0.65; 10: 0.68183; 11: 0.61667; 12: 0.3; 13: 0.675
| Median |  | 01: 0.56702; 03: 0.7; 04: 0.5; 05: 0.6; 06: 0.4; 07: 0.7; 09: 0.65; 10: 0.6855; 11: 0.6; 12: 0.3; 13: 0.66
| Std Dev. |  | 01: 0.089381; 03: 0.05164; 04: 0.20412; 05: 0; 07: 0.040825; 09: 0.054772; 10: 0.062017; 11: 0.05164; 12: 0; 13: 0.045935
| Rel.Std.Dev. |  | 01: 16.51%; 03: 7.746%; 04: 34.99%; 05: 0%; 07: 5.974%; 09: 8.427%; 10: 9.096%; 11: 8.374%; 12: 0%; 13: 6.805%
| PDM3 |  | 01: -12.96%; 03: 7.184%; 04: -6.214%; 05: -3.534%; 06: -35.69%; 07: 9.864%; 09: 4.504%; 10: 9.623%; 11: -0.8547%; 12: -51.77%; 13: 8.524%
| Z-Score (Absolute) |  | 01: 0.78; 03: 0.52; 04: 1.21; 05: 0.17; 06: 2.24; 07: 0.69; 09: 0.34; 10: 0.67; 11: 0; 12: 3.28; 13: 0.43
| Analytical results for K in OREAS 905 (Certified Value 2.94 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 19 | 20 | 22 | 30 | 31 | 32 | 33 |  |  |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  |  |  | 2
3
| 1 | 1 | 2.8 | 3.0516 | 2.78 | 2.9 | 3.2 | 2.9 | 3.03 | 3.01 | 2.856 | 3.1 | 3.02 | 2.48 | 3 | 3.176 | 2.8752 | 2.127 | 3.53 | 2.972 | 2.93 | 3.19 |  |  |  | 1
| 1 | 2 | 2.8 | 3.025 | 2.73 | 3 | 3.1 | 2.9 | 3.05 | 2.84 | 2.847 | 3.1 | 2.99 | 2.93 | 2.9 | 3.083 | 2.8514 | 1.947 | 3.64 | 3.013 | 2.92 | 3.07 |  |  |  | 0
| 1 | 3 | 2.7 | 3.0489 | 2.76 | 3 | 3.1 | 2.9 | 3.02 | 2.86 | 2.889 | 3 | 3.03 | 2.4 | 3 | 3.045 | 2.8534 | 2.016 | 3.77 | 3.013 | 2.93 | 2.91 |  |  |  | 16
| 1 | 4 | 2.7 | 2.8384 | 2.68 | 3 | 3.1 | 2.9 | 3 | 2.82 | 2.864 | 3 | 3.02 | 2.9 | 3 | 2.886 | 2.8231 | 2.016 | 3.57 | 3.08 | 2.97 | 2.73 |  |  |  | 2.9394
| 1 | 5 | 2.8 | 2.7168 | 2.76 | 3 | 3.1 | 3 | 3.07 | 2.83 | 2.823 | 3 | 3.01 | 2.67 | 3 | 2.964 | 2.8316 | 2.04 | 3.54 | 3.03 | 3.02 | 2.92 |  |  |  | 87
| 1 | 6 | 2.7 | 2.7995 | 2.73 | 3.1 | 3.2 | 2.9 | 3.01 | 2.83 | 2.897 | 3 | 3.04 | 3.04 | 3 | 2.992 | 2.8502 | 1.989 | 3.64 | 2.947 | 2.95 | 3.02
| Mean |  | 2.75 | 2.9134 | 2.74 | 3 | 3.1333 | 2.9167 | 3.03 | 2.865 | 2.8627 | 3.0333 | 3.0183 | 2.7367 | 2.9833 | 3.0243 | 2.8475 | 2.0225 | 3.615 | 3.0092 | 2.9533 | 2.9733
| Median |  | 2.75 | 2.9317 | 2.745 | 3 | 3.1 | 2.9 | 3.025 | 2.835 | 2.86 | 3 | 3.02 | 2.785 | 3 | 3.0185 | 2.8508 | 2.016 | 3.605 | 3.013 | 2.94 | 2.97
| Std Dev. |  | 0.054772 | 0.14636 | 0.035214 | 0.063246 | 0.05164 | 0.040825 | 0.026077 | 0.072319 | 0.02734 | 0.05164 | 0.017224 | 0.26067 | 0.040825 | 0.10069 | 0.01829 | 0.060222 | 0.08961 | 0.046318 | 0.037238 | 0.15782
| Rel.Std.Dev. |  | 1.992% | 5.024% | 1.285% | 2.108% | 1.648% | 1.4% | 0.8606% | 2.524% | 0.955% | 1.702% | 0.5706% | 9.525% | 1.368% | 3.329% | 0.6423% | 2.978% | 2.479% | 1.539% | 1.261% | 5.308%
| PDM3 |  | -6.445% | -0.8872% | -6.785% | 2.061% | 6.597% | -0.7745% | 3.081% | -2.532% | -2.612% | 3.195% | 2.684% | -6.898% | 1.494% | 2.888% | -3.128% | -31.19% | 22.98% | 2.372% | 0.4729% | 1.153%
| Z-Score (Absolute) |  | 2.11 | 0.49 | 2.2 | 0.36 | 1.68 | 0.46 | 0.66 | 1.26 | 0.99 | 0.69 | 0.54 | 0.77 | 0.2 | 0.6 | 1.14 | 9.28 | 6.43 | 0.45 | 0.1 | 0.1
| Analytical results for La in OREAS 905 (Certified Value 48 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 04 | 05 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 19 | 20 | 22 | 26 | 30 | 32 | [8 empty] | ppm
|  |  | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [8 empty] | 1
2
| 1 | 1 | 43 | 47.7 | 46 | 48.1 | 46 | 50.3 | 46.3 | 36 | 50 | 52.1 | 47.72 | 42 | 50.8 | 60 | 48.4 | [8 empty] | 1
| 1 | 2 | 43 | 48.4 | 47 | 47.6 | 47 | 48.9 | 47.1 | 42 | 46.9 | 52.3 | 49.175 | 37 | 50.3 | 40 | 48.1 | [8 empty] | 0
| 1 | 3 | 41 | 45.7 | 42 | 49.7 | 45 | 47.7 | 46.2 | 35 | 49 | 51.8 | 47.666 | 38 | 51.8 | 50 | 47.6 | [8 empty] | 16
| 1 | 4 | 38 | 49.8 | 48 | 50 | 47 | 47.9 | 47.5 | 39 | 50.6 | 50 | 47.11 | 38 | 48.2 | 50 | 46.4 | [8 empty] | 47.9702
| 1 | 5 | 45 | 46.7 | 48 | 50.5 | 45 | 43.3 | 46.5 | 37 | 50.3 | 51.5 | 45.828 | 37 | 52.4 | 50 | 48.7 | [8 empty] | 88
| 1 | 6 | 43 | 46.7 | 51 | 50.9 | 48 | 53 | 45.9 | 39 | 50.3 | 50.2 | 46.236 | 37 | 50 | 50 | 48.7
| Mean |  | 42.1667 | 47.5 | 47 | 49.4667 | 46.3333 | 48.5167 | 46.5833 | 38 | 49.5167 | 51.3167 | 47.2892 | 38.1667 | 50.5833 | 50 | 47.9833
| Median |  | 43 | 47.2 | 47.5 | 49.85 | 46.5 | 48.4 | 46.4 | 38 | 50.15 | 51.65 | 47.388 | 37.5 | 50.55 | 50 | 48.25
| Std Dev. |  | 2.4014 | 1.4601 | 2.9665 | 1.3277 | 1.2111 | 3.2152 | 0.60139 | 2.5298 | 1.3963 | 0.98268 | 1.1967 | 1.9408 | 1.4784 | 6.3246 | 0.87958
| Rel.Std.Dev. |  | 5.695% | 3.074% | 6.312% | 2.684% | 2.614% | 6.627% | 1.291% | 6.657% | 2.82% | 1.915% | 2.531% | 5.085% | 2.923% | 12.65% | 1.833%
| PDM3 |  | -12.1% | -0.9802% | -2.022% | 3.12% | -3.412% | 1.139% | -2.891% | -20.78% | 3.224% | 6.976% | -1.42% | -20.44% | 5.447% | 4.231% | 0.02739%
| Z-Score (Absolute) |  | 2.6 | 0 | 0.29 | 1.14 | 0.67 | 0.59 | 0.53 | 5.49 | 1.47 | 2.21 | 0.12 | 5.84 | 1.78 | 0.29 | 0.28
| Analytical results for Li in OREAS 905 (Certified Value 20.6 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 03 | 04 | 05 | 10 | 12 | 14 | 15 | 17 | 19 | 22 | 30 | 32 | 33 | [10 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES/MS | PF*OES | [10 empty] | 1
1
| 1 | 1 | 20.8432 | < 10 | 20 | 10 | 18 | 21 | 20 | 21 | 27.9 | 29 | 17.1 | 21.7 | 20 | [10 empty] | 1
| 1 | 2 | 20.1481 | < 10 | 20 | 10 | 18 | 21 | 20 | 21 | 24.6 | 25 | 16.5 | 20.7 | 20 | [10 empty] | 0
| 1 | 3 | 19.6495 | < 10 | 15 | 20 | 18 | 17 | 20 | 21 | 28 | 28 | 15.2 | 20.6 | 20 | [10 empty] | 16
| 1 | 4 | 20.7039 | < 10 | 20 | 10 | 18 | 20 | 20 | 22 | 24.2 | 24 | 16.8 | 21.9 | 20 | [10 empty] | 20.6086
| 1 | 5 | 20.4791 | < 10 | 15 | 10 | 16 | 21 | 20 | 22 | 26.8 | 23 | 16.5 | 22 | 20 | [10 empty] | 89
| 1 | 6 | 20.6409 | < 10 | 15 | 10 | 18 | 21 | 20 | 21 | 24.1 | 21 | 17.3 | 22 | 20
| Mean |  | 20.4108 |  | 17.5 | 11.6667 | 17.6667 | 20.1667 | 20 | 21.3333 | 25.9333 | 25 | 16.5667 | 21.4833 | 20
| Median |  | 20.56 |  | 17.5 | 10 | 18 | 21 | 20 | 21 | 25.7 | 24.5 | 16.65 | 21.8 | 20
| Std Dev. |  | 0.44248 |  | 2.7386 | 4.0825 | 0.8165 | 1.6021 | 0 | 0.5164 | 1.8457 | 3.0332 | 0.74207 | 0.65549 | 0
| Rel.Std.Dev. |  | 2.168% |  | 15.65% | 34.99% | 4.622% | 7.944% | 0% | 2.421% | 7.117% | 12.13% | 4.479% | 3.051% | 0%
| PDM3 |  | -0.9597% |  | -15.08% | -43.39% | -14.28% | -2.144% | -2.953% | 3.517% | 25.84% | 21.31% | -19.61% | 4.245% | -2.953%
| Z-Score (Absolute) |  | 0.07 | NA | 0.96 | 3.02 | 0.9 | 0.21 | 0.07 | 0.4 | 2.02 | 1.69 | 1.29 | 0.45 | 0.07
| Analytical results for Lu in OREAS 905 (Indicative Value 0.084 ppm) | [24 empty] | Indicative
Lu | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab |  |  |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 13; 06: 15; 07: 20; 09: 26; 10: 32 |  |  |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: BF*MS; 06: PF*OES/MS; 07: PF*MS; 09: ABL*MS; 10: PF*OES/MS |  |  |  | 3
3
| 1 | 1 | 01: < 0.25; 03: < 0.5; 04: < 0.05; 05: < 0.1; 06: 0.08; 07: < 0.1; 09: 0.11; 10: 0.106 |  |  |  | 1
| 1 | 2 | 01: < 0.25; 03: < 0.5; 04: < 0.05; 05: < 0.1; 06: 0.1; 07: < 0.1; 09: 0.1; 10: 0.096 |  |  |  | 6
| 1 | 3 | 01: < 0.25; 03: < 0.5; 04: < 0.05; 05: < 0.1; 06: 0.08; 07: < 0.1; 09: 0.08; 10: 0.102 |  |  |  | 16
| 1 | 4 | 01: < 0.25; 03: < 0.5; 04: < 0.05; 05: < 0.1; 06: 0.15; 07: 0.103; 09: 0.09; 10: 0.095 |  |  |  | 0.083733
| 1 | 5 | 01: < 0.25; 03: < 0.5; 04: < 0.05; 05: < 0.1; 06: 0.09; 07: 0.127; 09: 0.07; 10: 0.101 |  |  |  | 12
| 1 | 6 | 01: < 0.25; 03: < 0.5; 04: < 0.05; 05: < 0.1; 06: 0.08; 07: < 0.1; 09: 0.08; 10: 0.1
| Mean |  | 06: 0.096667; 07: 0.115; 09: 0.088333; 10: 0.1
| Median |  | 06: 0.085; 07: 0.115; 09: 0.085; 10: 0.1005
| Std Dev. |  | 06: 0.027325; 07: 0.016971; 09: 0.01472; 10: 0.0040497
| Rel.Std.Dev. |  | 06: 28.27%; 07: 14.76%; 09: 16.66%; 10: 4.05%
| PDM3 |  | 06: 15.45%; 07: 37.34%; 09: 5.494%; 10: 19.43%
| Z-Score (Absolute) |  | 01: NA; 03: NA; 04: NA; 05: NA; 06: 0.11; 07: 2.56; 09: 0.11; 10: 1.24
| Analytical results for Mg in OREAS 905 (Certified Value 0.288 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 20 | 22 | 26 | 30 | 31 | 32 | 33 | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*AAS | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | 3
3
| 1 | 1 | 0.277 | 0.27575 | 0.29 | 0.32 | 0.277 | 0.277 | 0.285 | 0.32 | 0.289 | 0.29 | 0.29 | 0.25 | 0.29 | 0.28 | 0.34 | 0.34 | 0.3079 | 0.314 | 0.31 | 0.44 | 0.271 | 0.286 | 0.27 | 1
| 1 | 2 | 0.277 | 0.26227 | 0.29 | 0.33 | 0.277 | 0.283 | 0.295 | 0.31 | 0.296 | 0.29 | 0.29 | 0.29 | 0.28 | 0.28 | 0.35 | 0.345 | 0.30794 | 0.2829 | 0.31 | 0.45 | 0.271 | 0.284 | 0.27 | 0
| 1 | 3 | 0.277 | 0.28213 | 0.29 | 0.32 | 0.271 | 0.283 | 0.285 | 0.32 | 0.296 | 0.28 | 0.29 | 0.24 | 0.29 | 0.28 | 0.34 | 0.344 | 0.30869 | 0.3036 | 0.29 | 0.45 | 0.271 | 0.278 | 0.28 | 16
| 1 | 4 | 0.277 | 0.27393 | 0.28 | 0.3 | 0.271 | 0.289 | 0.29 | 0.31 | 0.296 | 0.28 | 0.29 | 0.27 | 0.29 | 0.28 | 0.37 | 0.359 | 0.30843 | 0.3197 | 0.31 | 0.46 | 0.271 | 0.279 | 0.27 | 0.28787
| 1 | 5 | 0.277 | 0.25654 | 0.27 | 0.31 | 0.283 | 0.289 | 0.295 | 0.31 | 0.289 | 0.28 | 0.29 | 0.26 | 0.29 | 0.28 | 0.32 | 0.321 | 0.30588 | 0.3071 | 0.29 | 0.42 | 0.271 | 0.277 | 0.3 | 90
| 1 | 6 | 0.277 | 0.26288 | 0.29 | 0.3 | 0.271 | 0.289 | 0.29 | 0.29 | 0.296 | 0.28 | 0.29 | 0.28 | 0.28 | 0.28 | 0.33 | 0.327 | 0.31203 | 0.3048 | 0.31 | 0.42 | 0.271 | 0.285 | 0.28
| Mean |  | 0.277 | 0.26892 | 0.285 | 0.31333 | 0.275 | 0.285 | 0.29 | 0.31 | 0.29367 | 0.28333 | 0.29 | 0.265 | 0.28667 | 0.28 | 0.34167 | 0.33933 | 0.30848 | 0.30535 | 0.30333 | 0.44 | 0.271 | 0.2815 | 0.27833
| Median |  | 0.277 | 0.2684 | 0.29 | 0.315 | 0.274 | 0.286 | 0.29 | 0.31 | 0.296 | 0.28 | 0.29 | 0.265 | 0.29 | 0.28 | 0.34 | 0.342 | 0.30818 | 0.30595 | 0.31 | 0.445 | 0.271 | 0.2815 | 0.275
| Std Dev. |  | 6.0809e-17 | 0.0097997 | 0.0083666 | 0.012111 | 0.004899 | 0.004899 | 0.0044721 | 0.010954 | 0.0036148 | 0.005164 | 0 | 0.018708 | 0.005164 | 0 | 0.017224 | 0.013633 | 0.0020032 | 0.012579 | 0.010328 | 0.016733 | 6.0809e-17 | 0.003937 | 0.01169
| Rel.Std.Dev. |  | 2.195e-14% | 3.644% | 2.936% | 3.865% | 1.781% | 1.719% | 1.542% | 3.534% | 1.231% | 1.823% | 0% | 7.06% | 1.801% | 0% | 5.041% | 4.018% | 0.6494% | 4.12% | 3.405% | 3.803% | 2.244e-14% | 1.399% | 4.2%
| PDM3 |  | -3.776% | -6.584% | -0.9965% | 8.846% | -4.47% | -0.9965% | 0.7404% | 7.688% | 2.014% | -1.575% | 0.7404% | -7.944% | -0.4175% | -2.733% | 18.69% | 17.88% | 7.16% | 6.073% | 5.372% | 52.85% | -5.86% | -2.212% | -3.312%
| Z-Score (Absolute) |  | 0.55 | 1.06 | 0.1 | 1.6 | 0.67 | 0.07 | 0.2 | 1.4 | 0.41 | 0.2 | 0.2 | 1 | 0 | 0.4 | 3.29 | 3.15 | 1.31 | 1.12 | 1 | 9.18 | 0.91 | 0.31 | 0.5
| Analytical results for Mn in OREAS 905 (Certified Value 0.039 wt.%) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 26 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*MS | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*AAS | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.023 | 0.037865 | 0.0383 | 0.0395 | 0.046 | 0.039 | 0.04 | 0.04 | 0.039 | 0.0411 | 0.04 | 0.0363 | 0.04 | 0.0403 | 0.03 | 0.0399 | 0.0322 | 0.0387 | 0.039 | 0.046 | 0.039 | 0.038 |  | 1
| 1 | 2 | 0.023 | 0.038102 | 0.0391 | 0.041 | 0.039 | 0.039 | 0.04 | 0.04 | 0.039 | 0.0429 | 0.04 | 0.041 | 0.04 | 0.0406 | 0.03 | 0.0408 | 0.0299 | 0.0426 | 0.04 | 0.039 | 0.0395 | 0.035 |  | 0
| 1 | 3 | 0.023 | 0.039262 | 0.039 | 0.04 | 0.039 | 0.039 | 0.04 | 0.04 | 0.039 | 0.0407 | 0.04 | 0.0357 | 0.04 | 0.0395 | 0.03 | 0.0402 | 0.032056 | 0.0407 | 0.039 | 0.039 | 0.0376 | 0.036 |  | 16
| 1 | 4 | 0.023 | 0.038754 | 0.038 | 0.04 | 0.039 | 0.039 | 0.04 | 0.04 | 0.039 | 0.0415 | 0.04 | 0.04 | 0.04 | 0.0397 | 0.03 | 0.0403 | 0.0345 | 0.0409 | 0.04 | 0.039 | 0.0388 | 0.034 |  | 0.039483
| 1 | 5 | 0.023 | 0.037123 | 0.0378 | 0.0395 | 0.039 | 0.039 | 0.04 | 0.04 | 0.039 | 0.0393 | 0.04 | 0.0377 | 0.04 | 0.0399 | 0.03 | 0.0407 | 0.031913 | 0.0433 | 0.039 | 0.039 | 0.0386 | 0.035 |  | 91
| 1 | 6 | 0.023 | 0.037803 | 0.0383 | 0.037 | 0.039 | 0.039 | 0.04 | 0.04 | 0.039 | 0.044 | 0.04 | 0.041 | 0.039 | 0.0405 | 0.03 | 0.0406 | 0.031913 | 0.0388 | 0.038 | 0.039 | 0.0387 | 0.035
| Mean |  | 0.023 | 0.038151 | 0.038417 | 0.0395 | 0.040167 | 0.039 | 0.04 | 0.04 | 0.039 | 0.041583 | 0.04 | 0.038617 | 0.039833 | 0.040083 | 0.03 | 0.040417 | 0.03208 | 0.040833 | 0.039167 | 0.040167 | 0.0387 | 0.0355
| Median |  | 0.023 | 0.037983 | 0.0383 | 0.03975 | 0.039 | 0.039 | 0.04 | 0.04 | 0.039 | 0.0413 | 0.04 | 0.03885 | 0.04 | 0.0401 | 0.03 | 0.04045 | 0.031984 | 0.0408 | 0.039 | 0.039 | 0.03875 | 0.035
| Std Dev. |  | 3.8006e-18 | 0.00075641 | 0.00052694 | 0.0013416 | 0.0028577 | 0 | 0 | 0 | 0 | 0.0016618 | 0 | 0.0023659 | 0.00040825 | 0.00044907 | 0 | 0.00034303 | 0.0014615 | 0.0018928 | 0.00075277 | 0.0028577 | 0.0006261 | 0.0013784
| Rel.Std.Dev. |  | 1.652e-14% | 1.983% | 1.372% | 3.397% | 7.115% | 0% | 0% | 0% | 0% | 3.996% | 0% | 6.127% | 1.025% | 1.12% | 0% | 0.8487% | 4.556% | 4.635% | 1.922% | 7.115% | 1.618% | 3.883%
| PDM3 |  | -41.75% | -3.373% | -2.701% | 0.04276% | 1.731% | -1.224% | 1.309% | 1.309% | -1.224% | 5.319% | 1.309% | -2.194% | 0.887% | 1.52% | -24.02% | 2.364% | -18.75% | 3.42% | -0.8015% | 1.731% | -1.983% | -10.09%
| Z-Score (Absolute) |  | 8.18 | 0.3 | 0.16 | 0.67 | 0 | 0 | 0.67 | 0.67 | 0 | 1.51 | 0.67 | 0.06 | 0.59 | 0.72 | 4.61 | 0.89 | 3.51 | 1.11 | 0.23 | 0 | 0.01 | 1.97
| Analytical results for Mo in OREAS 905 (Indicative Value 2.95 ppm) | [24 empty] | Indicative
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 32 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 2.8852 | < 10 | 5.1 | 2 | 4 | 2 | 3 | < 10 | 3 | 3 | 9 | < 5 | 2.4 | 3.1 | [9 empty] | 1
| 1 | 2 | 2.6403 | < 10 | 4.1 | 2 | 3 | 1 | 3 | < 10 | 3 | 3 | 8.6 | < 5 | 2.4 | 3.7 | [9 empty] | 7
| 1 | 3 | 3.1129 | < 10 | 4.8 | 2 | 4 | 2 | 2 | < 10 | 3 | 2 | 9.4 | < 5 | 3.4 | 3.5 | [9 empty] | 16
| 1 | 4 | 4.253 | < 10 | 4.7 | 4 | 5 | 2 | 2 | < 10 | 3 | 3 | 9.3 | < 5 | 2.6 | 3.3 | [9 empty] | 2.9487
| 1 | 5 | 3.1673 | < 10 | 5.3 | 4 | 4 | 1 | 2 | < 10 | 3 | 3 | 9.7 | < 5 | 3.2 | 5.2 | [9 empty] | 13
| 1 | 6 | 3.3836 | < 10 | 5.1 | 4 | 5 | 4 | 2 | < 10 | 3 | 3 | 9.5 | < 5 | 2.6 | 3.4
| Mean |  | 01: 3.2404; 04: 4.85; 05: 3; 06: 4.1667; 07: 2; 09: 2.3333; 11: 3; 12: 2.8333; 13: 9.25; 15: 2.7667; 17: 3.7
| Median |  | 01: 3.1401; 04: 4.95; 05: 3; 06: 4; 07: 2; 09: 2; 11: 3; 12: 3; 13: 9.35; 15: 2.6; 17: 3.45
| Std Dev. |  | 01: 0.55745; 04: 0.42778; 05: 1.0954; 06: 0.75277; 07: 1.0954; 09: 0.5164; 11: 0; 12: 0.40825; 13: 0.3937; 15: 0.4274; 17: 0.76158
| Rel.Std.Dev. |  | 01: 17.2%; 04: 8.82%; 05: 36.51%; 06: 18.07%; 07: 54.77%; 09: 22.13%; 11: 0%; 12: 14.41%; 13: 4.256%; 15: 15.45%; 17: 20.58%
| PDM3 |  | 01: 9.894%; 04: 64.48%; 05: 1.741%; 06: 41.31%; 07: -32.17%; 09: -20.87%; 11: 1.741%; 12: -3.911%; 13: 213.7%; 15: -6.172%; 17: 25.48%
| Z-Score (Absolute) |  | 0.06 | NA | 3.12 | 0 | 1.97 | 1.69 | 1.12 | NA | 0 | 0.28 | 10.54 | NA | 0.39 | 0.67
| Analytical results for Na in OREAS 905 (Indicative Value 2.41 wt.%) | [24 empty] | Indicative
Na | Round | Replicate | 01: Lab; 03: Lab; 04: Lab |  | 1
| No | No | 01: 09; 03: 11; 04: 13 |  | wt.%
|  |  | 01: BF*OES; 03: BF*OES; 04: BF*OES |  | 2
2
| 1 | 1 | 01: 2.42; 03: 2.307; 04: 2.41 |  | 1
| 1 | 2 | 01: 2.4; 03: 2.3; 04: 2.43 |  | 8
| 1 | 3 | 01: 2.47; 03: 2.329; 04: 2.45 |  | 16
| 1 | 4 | 01: 2.49; 03: 2.329; 04: 2.4 |  | 2.4075
| 1 | 5 | 01: 2.55; 03: 2.315; 04: 2.42 |  | 14
| 1 | 6 | 01: 2.48; 03: 2.344; 04: 2.49
| Mean |  | 01: 2.4683; 03: 2.3207; 04: 2.4333
| Median |  | 01: 2.475; 03: 2.322; 04: 2.425
| Std Dev. |  | 01: 0.053448; 03: 0.016305; 04: 0.03266
| Rel.Std.Dev. |  | 01: 2.165%; 03: 0.7026%; 04: 1.342%
| PDM3 |  | 01: 2.528%; 03: -3.605%; 04: 1.075%
| Z-Score (Absolute) |  | 01: 0.67; 03: 2.17; 04: 0
| Analytical results for Nb in OREAS 905 (Certified Value 18.5 ppm) | [24 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 05 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 26 | 30 | 32 | [9 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | ABL*MS | PF*MS | PF*OES/MS | [9 empty] | 1
1
| 1 | 1 | 18.7279 | 25 | 15 | 13 | 20 | 19 | 44 | 19 | 16 | 8 | 21.8856 | 18.6 | 30 | 19.4 | [9 empty] | 1
| 1 | 2 | 19.4443 | 25 | 15 | 15 | 17.3 | 19.5 | 58 | 15 | 16 | 8.5 | 17.7529 | 19.2 | 31 | 19.3 | [9 empty] | 15
| 1 | 3 | 19.8182 | 20 | 15 | 16 | 15.9 | 18.9 | 11 | 19 | 17 | 8.3 | 21.9505 | 19.1 | 30 | 19.1 | [9 empty] | 16
| 1 | 4 | 19.9538 | 25 | 15 | 16 | 16.6 | 18.6 | 10 | 19 | 19 | 8.2 | 17.5329 | 18.1 | 31 | 18.6 | [9 empty] | 18.5042
| 1 | 5 | 18.5712 | 25 | 20 | 14 | 13.6 | 18.8 | 43 | 17 | 16 | 9 | 21.0342 | 19.5 | 28 | 19.2 | [9 empty] | 92
| 1 | 6 | 17.2852 | 25 | 20 | 16 | 26.1 | 18.9 | < 10 | 18 | 18 | 7.5 | 20.6437 | 19.4 | 28 | 19.4
| Mean |  | 18.9668 | 24.1667 | 16.6667 | 15 | 18.25 | 18.95 | 33.2 | 17.8333 | 17 | 8.25 | 20.1333 | 18.9833 | 29.6667 | 19.1667
| Median |  | 19.0861 | 25 | 15 | 15.5 | 16.95 | 18.9 | 43 | 18.5 | 16.5 | 8.25 | 20.839 | 19.15 | 30 | 19.25
| Std Dev. |  | 0.99684 | 2.0412 | 2.582 | 1.2649 | 4.3693 | 0.30166 | 21.5569 | 1.6021 | 1.2649 | 0.501 | 1.9937 | 0.53448 | 1.3663 | 0.30111
| Rel.Std.Dev. |  | 5.256% | 8.447% | 15.49% | 8.433% | 23.94% | 1.592% | 64.93% | 8.984% | 7.441% | 6.073% | 9.902% | 2.816% | 4.605% | 1.571%
| PDM3 |  | 2.5% | 30.6% | -9.931% | -18.94% | -1.374% | 2.409% | 79.42% | -3.626% | -8.129% | -55.42% | 8.804% | 2.589% | 60.32% | 3.58%
| Z-Score (Absolute) |  | 0 | 1.66 | 0.73 | 1.26 | 0.73 | 0 | 4.53 | 0.36 | 0.62 | 3.41 | 0.37 | 0.01 | 3.41 | 0.07
| Analytical results for Nd in OREAS 905 (Certified Value 40.5 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 30; 12: 32 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*MS; 12: PF*OES/MS |  | 1
2
| 1 | 1 | 01: 39.2; 03: 40; 04: 40; 05: 39.2; 06: 39.5; 07: 41.3; 09: 40.588; 10: 41.1; 11: 48.9; 12: 41.9 |  | 1
| 1 | 2 | 01: 37.4; 03: 42.5; 04: 39.8; 05: 38.5; 06: 39.6; 07: 39.5; 09: 37.932; 10: 43.6; 11: 48.2; 12: 40.5 |  | 0
| 1 | 3 | 01: 38.4; 03: 41.5; 04: 41.1; 05: 37.1; 06: 38.4; 07: 39.9; 09: 39.498; 10: 43.2; 11: 48.7; 12: 40.6 |  | 16
| 1 | 4 | 01: 41.5; 03: 40.5; 04: 42.7; 05: 37.6; 06: 39.2; 07: 41.9; 09: 38.321; 10: 41.3; 11: 46.1; 12: 39.3 |  | 40.4846
| 1 | 5 | 01: 39; 03: 45; 04: 42; 05: 40.4; 06: 40.3; 07: 41.2; 09: 39.842; 10: 43.9; 11: 47; 12: 41.8 |  | 93
| 1 | 6 | 01: 37.9; 03: 45; 04: 43.1; 05: 40.4; 06: 40.7; 07: 41.5; 09: 39.09; 10: 40; 11: 49.4; 12: 40.9
| Mean |  | 01: 38.9; 03: 42.4167; 04: 41.45; 05: 38.8667; 06: 39.6167; 07: 40.8833; 09: 39.2118; 10: 42.1833; 11: 48.05; 12: 40.8333
| Median |  | 01: 38.7; 03: 42; 04: 41.55; 05: 38.85; 06: 39.55; 07: 41.25; 09: 39.294; 10: 42.25; 11: 48.45; 12: 40.75
| Std Dev. |  | 01: 1.4394; 03: 2.1775; 04: 1.3809; 05: 1.3909; 06: 0.81343; 07: 0.95586; 09: 0.98176; 10: 1.5943; 11: 1.2566; 12: 0.95847
| Rel.Std.Dev. |  | 01: 3.7%; 03: 5.134%; 04: 3.332%; 05: 3.579%; 06: 2.053%; 07: 2.338%; 09: 2.504%; 10: 3.779%; 11: 2.615%; 12: 2.347%
| PDM3 |  | 01: -3.914%; 03: 4.772%; 04: 2.384%; 05: -3.997%; 06: -2.144%; 07: 0.9848%; 09: -3.144%; 10: 4.196%; 11: 18.69%; 12: 0.8613%
| Z-Score (Absolute) |  | 01: 0.92; 03: 0.73; 04: 0.28; 05: 0.93; 06: 0.58; 07: 0.01; 09: 0.77; 10: 0.62; 11: 3.36; 12: 0.01
| Analytical results for Ni in OREAS 905 (Indicative Value 12.5 ppm) | [24 empty] | Indicative
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 06 | 10 | 11 | 12 | 14 | 15 | 19 | 22 | 26 | 30 | 32 | 33 | [8 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | PF*OES | PF*OES/MS | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | PF*OES | [8 empty] | 1
1
| 1 | 1 | < 20 | < 10 | < 10 | < 50 | 6 | < 50 | 10 | < 10 | 9 | 13.5 | 18 | 14 | 40 | 8.8 | < 20 | [8 empty] | 1
| 1 | 2 | < 20 | < 10 | < 10 | < 50 | 15 | < 50 | 20 | < 10 | 6 | 13.2 | 16 | 16 | 40 | 11.3 | < 20 | [8 empty] | 9
| 1 | 3 | < 20 | < 10 | < 10 | < 50 | 9 | < 50 | 30 | < 10 | 9 | 13.9 | 19 | 14 | 40 | 9.3 | < 20 | [8 empty] | 16
| 1 | 4 | < 20 | < 10 | < 10 | < 50 | 7 | < 50 | 20 | < 10 | 8 | 13.5 | 20 | 16 | 50 | 7.7 | < 20 | [8 empty] | 12.5381
| 1 | 5 | < 20 | < 10 | < 10 | < 50 | 7 | < 50 | < 10 | < 10 | 11 | 13 | 16 | 16 | 40 | 6.9 | < 20 | [8 empty] | 15
| 1 | 6 | < 20 | < 10 | < 10 | < 50 | 8 | < 50 | 20 | < 10 | 7 | 13.3 | 17 | 12 | 40 | 31.4 | < 20
| Mean |  | 06: 8.6667; 09: 20; 11: 8.3333; 12: 13.4; 13: 17.6667; 14: 14.6667; 15: 41.6667; 17: 12.5667
| Median |  | 06: 7.5; 09: 20; 11: 8.5; 12: 13.4; 13: 17.5; 14: 15; 15: 40; 17: 9.05
| Std Dev. |  | 06: 3.266; 09: 7.0711; 11: 1.7512; 12: 0.30984; 13: 1.633; 14: 1.633; 15: 4.0825; 17: 9.3483
| Rel.Std.Dev. |  | 06: 37.68%; 09: 35.36%; 11: 21.01%; 12: 2.312%; 13: 9.243%; 14: 11.13%; 15: 9.798%; 17: 74.39%
| PDM3 |  | 06: -30.88%; 09: 59.51%; 11: -33.54%; 12: 6.874%; 13: 40.9%; 14: 16.98%; 15: 232.3%; 17: 0.2279%
| Z-Score (Absolute) |  | NA | NA | NA | NA | 1.01 | NA | 0.53 | NA | 0.87 | 0.1 | 0.55 | 0.1 | 4.2 | 0.8 | NA
| Analytical results for P in OREAS 905 (Certified Value 0.028 wt.%) | [24 empty] | Certified
P | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 04; 03: 09; 04: 10; 05: 11; 06: 12; 07: 13; 09: 14; 10: 15; 11: 19; 12: 22; 13: 30; 14: 32 |  | wt.%
|  |  | 01: PF*OES; 03: BF*OES; 04: PF*OES/MS; 05: BF*OES; 06: PF*OES; 07: BF*OES; 09: PF*OES; 10: PF*OES/MS; 11: PF*OES; 12: PF*OES/MS; 13: PF*OES; 14: PF*OES/MS |  | 3
3
| 1 | 1 | 01: 0.04; 03: 0.03; 04: 0.03; 05: 0.026; 06: 0.028; 07: 0.03; 09: 0.03; 10: 0.03; 11: 0.0283; 12: 0.0305; 13: < 0.01; 14: 0.027 |  | 1
| 1 | 2 | 01: 0.03; 03: 0.03; 04: 0.02; 05: 0.026; 06: 0.027; 07: 0.03; 09: 0.03; 10: 0.03; 11: 0.0285; 12: 0.0243; 13: < 0.01; 14: 0.0248 |  | 0
| 1 | 3 | 01: 0.04; 03: 0.03; 04: 0.03; 05: 0.026; 06: 0.026; 07: 0.03; 09: 0.03; 10: 0.03; 11: 0.0284; 12: 0.0297; 13: < 0.01; 14: 0.0276 |  | 16
| 1 | 4 | 01: 0.04; 03: 0.035; 04: 0.03; 05: 0.026; 06: 0.026; 07: 0.02; 09: 0.03; 10: 0.03; 11: 0.0276; 12: 0.0369; 13: < 0.01; 14: 0.0246 |  | 0.02829
| 1 | 5 | 01: 0.03; 03: 0.035; 04: 0.03; 05: 0.026; 06: 0.026; 07: 0.02; 09: 0.03; 10: 0.03; 11: 0.0289; 12: 0.0298; 13: < 0.01; 14: 0.0285 |  | 94
| 1 | 6 | 01: 0.03; 03: 0.03; 04: 0.03; 05: 0.026; 06: 0.026; 07: 0.03; 09: 0.03; 10: 0.03; 11: 0.0275; 12: 0.0298; 13: < 0.01; 14: 0.0267
| Mean |  | 01: 0.035; 03: 0.031667; 04: 0.028333; 05: 0.026; 06: 0.0265; 07: 0.026667; 09: 0.03; 10: 0.03; 11: 0.0282; 12: 0.030167; 14: 0.026533
| Median |  | 01: 0.035; 03: 0.03; 04: 0.03; 05: 0.026; 06: 0.026; 07: 0.03; 09: 0.03; 10: 0.03; 11: 0.02835; 12: 0.0298; 14: 0.02685
| Std Dev. |  | 01: 0.0054772; 03: 0.002582; 04: 0.0040825; 05: 0; 06: 0.00083666; 07: 0.005164; 09: 0; 10: 0; 11: 0.00054406; 12: 0.0040088; 14: 0.0015488
| Rel.Std.Dev. |  | 01: 15.65%; 03: 8.154%; 04: 14.41%; 05: 0%; 06: 3.157%; 07: 19.36%; 09: 0%; 10: 0%; 11: 1.929%; 12: 13.29%; 14: 5.837%
| PDM3 |  | 01: 23.72%; 03: 11.93%; 04: 0.1523%; 05: -8.096%; 06: -6.328%; 07: -5.739%; 09: 6.044%; 10: 6.044%; 11: -0.319%; 12: 6.633%; 14: -6.21%
| Z-Score (Absolute) |  | 01: 2.7; 03: 1.35; 04: 0; 05: 0.87; 06: 0.74; 07: 0.67; 09: 0.67; 10: 0.67; 11: 0.05; 12: 0.2; 13: NA; 14: 0.73
| Analytical results for Pb in OREAS 905 (Certified Value 30.8 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | 33 | [7 empty] | ppm
|  |  | PF*OES | PF*MS | PF*OES | PF*OES | PF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | PF*OES | [7 empty] | 1
1
| 1 | 1 | < 100 | 32.4519 | < 20 | 30 | < 100 | 28.1 | 33 | < 20 | 33 | 34 | 47.9 | 21 | 32 | 20 | 32.7 | < 300 | [7 empty] | 1
| 1 | 2 | < 100 | 36.6369 | < 20 | 25 | < 100 | 25.6 | 32 | 46 | 31 | 37 | 45.8 | 21 | 31 | 15 | 32 | < 300 | [7 empty] | 0
| 1 | 3 | < 100 | 34.2908 | < 20 | 30 | < 100 | 29.1 | 32 | 25 | 31 | 38 | 45 | 20 | 31 | 12 | 31.6 | < 300 | [7 empty] | 16
| 1 | 4 | < 100 | 32.083 | < 20 | 25 | < 100 | 26.1 | 28 | 32 | 33 | 37 | 47.1 | 26 | 30 | 9 | 31.3 | < 300 | [7 empty] | 30.8094
| 1 | 5 | < 100 | 29.2818 | < 20 | 30 | < 100 | 26.5 | 28 | 26 | 32 | 36 | 47.7 | 23 | 32 | 8 | 31.3 | < 300 | [7 empty] | 95
| 1 | 6 | < 100 | 31.5845 | < 20 | 30 | < 100 | 35.7 | 27 | < 20 | 32 | 38 | 46.9 | 20 | 31 | 10 | 31 | < 300
| Mean |  | 03: 32.7215; 05: 28.3333; 07: 28.5167; 09: 30; 10: 32.25; 11: 32; 12: 36.6667; 13: 46.7333; 14: 21.8333; 15: 31.1667; 17: 12.3333; 18: 31.65
| Median |  | 03: 32.2675; 05: 30; 07: 27.3; 09: 30; 10: 29; 11: 32; 12: 37; 13: 47; 14: 21; 15: 31; 17: 11; 18: 31.45
| Std Dev. |  | 03: 2.5059; 05: 2.582; 07: 3.7558; 09: 2.6077; 10: 9.6738; 11: 0.89443; 12: 1.5055; 13: 1.1255; 14: 2.3166; 15: 0.75277; 17: 4.5019; 18: 0.61563
| Rel.Std.Dev. |  | 03: 7.658%; 05: 9.113%; 07: 13.17%; 09: 8.692%; 10: 30%; 11: 2.795%; 12: 4.106%; 13: 2.408%; 14: 10.61%; 15: 2.415%; 17: 36.5%; 18: 1.945%
| PDM3 |  | 03: 6.206%; 05: -8.037%; 07: -7.442%; 09: -2.627%; 10: 4.676%; 11: 3.864%; 12: 19.01%; 13: 51.69%; 14: -29.13%; 15: 1.16%; 17: -59.97%; 18: 2.728%
| Z-Score (Absolute) |  | NA | 0.56 | NA | 0.59 | NA | 0.91 | 0.15 | 0.76 | 0.37 | 1.59 | 4.22 | 2.5 | 0.15 | 4.76 | 0.28 | NA
| Analytical results for Pr in OREAS 905 (Certified Value 10.9 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 30; 12: 32 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*MS; 12: PF*OES/MS |  | 1
2
| 1 | 1 | 01: 9.8; 03: 11; 04: 10.3; 05: 10.2; 06: 10.5; 07: 11.4; 09: 10.609; 10: 11.6; 11: 14.5; 12: 11.3 |  | 1
| 1 | 2 | 01: 10.1; 03: 12; 04: 10.3; 05: 9.8; 06: 10.8; 07: 10.4; 09: 11.069; 10: 11.8; 11: 14.6; 12: 11.2 |  | 0
| 1 | 3 | 01: 9.7; 03: 10; 04: 10.8; 05: 9.8; 06: 10.7; 07: 11.1; 09: 10.424; 10: 11.8; 11: 15.5; 12: 11.1 |  | 16
| 1 | 4 | 01: 10.5; 03: 12; 04: 11; 05: 9.8; 06: 10.8; 07: 11.6; 09: 10.714; 10: 11.3; 11: 14; 12: 10.8 |  | 10.8876
| 1 | 5 | 01: 9.7; 03: 13; 04: 10.9; 05: 8.8; 06: 10.6; 07: 11.4; 09: 10.844; 10: 12.2; 11: 14.3; 12: 11.1 |  | 96
| 1 | 6 | 01: 9.9; 03: 12; 04: 11.1; 05: 10.4; 06: 10.7; 07: 11.5; 09: 10.573; 10: 11.8; 11: 14.6; 12: 11.1
| Mean |  | 01: 9.95; 03: 11.6667; 04: 10.7333; 05: 9.8; 06: 10.6833; 07: 11.2333; 09: 10.7055; 10: 11.75; 11: 14.5833; 12: 11.1
| Median |  | 01: 9.85; 03: 12; 04: 10.85; 05: 9.8; 06: 10.7; 07: 11.4; 09: 10.6615; 10: 11.8; 11: 14.55; 12: 11.1
| Std Dev. |  | 01: 0.30822; 03: 1.0328; 04: 0.35024; 05: 0.55136; 06: 0.1169; 07: 0.44121; 09: 0.22694; 10: 0.29496; 11: 0.50365; 12: 0.16733
| Rel.Std.Dev. |  | 01: 3.098%; 03: 8.853%; 04: 3.263%; 05: 5.626%; 06: 1.094%; 07: 3.928%; 09: 2.12%; 10: 2.51%; 11: 3.454%; 12: 1.507%
| PDM3 |  | 01: -8.612%; 03: 7.155%; 04: -1.417%; 05: -9.99%; 06: -1.877%; 07: 3.175%; 09: -1.673%; 10: 7.92%; 11: 33.94%; 12: 1.95%
| Z-Score (Absolute) |  | 01: 1.06; 03: 0.82; 04: 0.2; 05: 1; 06: 0.26; 07: 0.53; 09: 0.23; 10: 0.91; 11: 4.01; 12: 0.2
| Analytical results for Rb in OREAS 905 (Certified Value 137 ppm) | [24 empty] | Certified
Rb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 09: 15; 10: 17; 11: 20; 12: 22; 13: 26; 14: 30; 15: 32 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: BF*MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*MS; 12: PF*OES/MS; 13: ABL*MS; 14: PF*MS; 15: PF*OES/MS |  | 0
0
| 1 | 1 | 01: 143.1227; 03: 141; 04: 153; 05: 134; 06: 142; 07: 137; 09: 138; 10: 135.7; 11: 130.087; 12: 202; 13: 138; 14: 93.6; 15: 144 |  | 1
| 1 | 2 | 01: 143.0064; 03: 140; 04: 151; 05: 130; 06: 144; 07: 140; 09: 135; 10: 138.5; 11: 131.5009; 12: 180; 13: 135; 14: 91.7; 15: 142 |  | 0
| 1 | 3 | 01: 142.7734; 03: 139; 04: 144; 05: 131; 06: 136; 07: 140; 09: 135; 10: 135.3; 11: 137.0306; 12: 190; 13: 140; 14: 92.5; 15: 143 |  | 16
| 1 | 4 | 01: 141.6073; 03: 142; 04: 158; 05: 130; 06: 140; 07: 141.5; 09: 138; 10: 133.1; 11: 130.8463; 12: 184; 13: 134; 14: 89; 15: 141 |  | 137.4463
| 1 | 5 | 01: 139.9391; 03: 136; 04: 156; 05: 129; 06: 130; 07: 138.6; 09: 136; 10: 134.2; 11: 133.325; 12: 192; 13: 145; 14: 88.6; 15: 142 |  | 97
| 1 | 6 | 01: 141.2696; 03: 136; 04: 155; 05: 130; 06: 145; 07: 139.3; 09: 138; 10: 133.5; 11: 134.5713; 12: 185; 13: 132; 14: 90.7; 15: 144
| Mean |  | 01: 141.9531; 03: 139; 04: 152.8333; 05: 130.6667; 06: 139.5; 07: 139.4; 09: 136.6667; 10: 135.05; 11: 132.8935; 12: 188.8333; 13: 137.3333; 14: 91.0167; 15: 142.6667
| Median |  | 01: 142.1904; 03: 139.5; 04: 154; 05: 130; 06: 141; 07: 139.65; 09: 137; 10: 134.75; 11: 132.413; 12: 187.5; 13: 136.5; 14: 91.2; 15: 142.5
| Std Dev. |  | 01: 1.2485; 03: 2.5298; 04: 4.9565; 05: 1.7512; 06: 5.648; 07: 1.5192; 09: 1.5055; 10: 1.9655; 11: 2.6124; 12: 7.7567; 13: 4.7188; 14: 1.9671; 15: 1.2111
| Rel.Std.Dev. |  | 01: 0.8795%; 03: 1.82%; 04: 3.243%; 05: 1.34%; 06: 4.049%; 07: 1.09%; 09: 1.102%; 10: 1.455%; 11: 1.966%; 12: 4.108%; 13: 3.436%; 14: 2.161%; 15: 0.8489%
| PDM3 |  | 01: 3.279%; 03: 1.13%; 04: 11.19%; 05: -4.933%; 06: 1.494%; 07: 1.421%; 09: -0.5672%; 10: -1.743%; 11: -3.312%; 12: 37.39%; 13: -0.08221%; 14: -33.78%; 15: 3.798%
| Z-Score (Absolute) |  | 01: 0.54; 03: 0; 04: 2.54; 05: 1.66; 06: 0.09; 07: 0.07; 09: 0.43; 10: 0.73; 11: 1.12; 12: 9.16; 13: 0.31; 14: 8.82; 15: 0.67
| Analytical results for Re in OREAS 905 (Indicative Value < 0.1 ppm) | [24 empty] | Indicative
Re | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 03; 03: 09; 04: 17; 05: 26; 06: 30 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*MS; 05: ABL*MS; 06: PF*MS |  | 2
2
| 1 | 1 | 01: < 0.1; 03: < 0.1; 04: < 0.1; 05: 0.01; 06: - |  | 1
| 1 | 2 | 01: < 0.1; 03: < 0.1; 04: < 0.1; 05: < 0.01; 06: - |  | 4
| 1 | 3 | 01: < 0.1; 03: < 0.1; 04: < 0.1; 05: 0.02; 06: - |  | 16
| 1 | 4 | 01: < 0.1; 03: < 0.1; 04: < 0.1; 05: < 0.01; 06: - |  | < 0.1
| 1 | 5 | 01: < 0.1; 03: < 0.1; 04: < 0.1; 05: 0.02; 06: - |  | 16
| 1 | 6 | 01: < 0.1; 03: < 0.1; 04: < 0.1; 05: < 0.01; 06: -
| Mean |  | 05: 0.016667
| Median |  | 05: 0.02
| Std Dev. |  | 05: 0.0057735
| Rel.Std.Dev. |  | 05: 34.64%
PDM3
Z-Score (Absolute) | bar: 01 to 06
| Analytical results for S in OREAS 905 (Certified Value 0.071 wt.%) | [24 empty] | Certified
S | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 01; 03: 03; 04: 05; 05: 06; 06: 07; 07: 12; 09: 13; 10: 17; 11: 19; 12: 22; 13: 31; 14: 33 |  | wt.%
|  |  | 01: PF*OES; 03: PF*OES; 04: PF*OES; 05: PF*OES; 06: PF*OES; 07: PF*OES; 09: BF*OES; 10: PF*OES; 11: PF*OES; 12: PF*OES/MS; 13: PF*OES; 14: PF*OES |  | 3
3
| 1 | 1 | 01: 0.07; 03: < 0.05; 04: 0.07; 05: 0.07; 06: 0.06; 07: 0.09; 09: 0.08; 10: 0.08; 11: 0.098; 12: 0.0608; 13: 0.02; 14: 0.05 |  | 1
| 1 | 2 | 01: 0.07; 03: < 0.05; 04: 0.06; 05: 0.09; 06: 0.06; 07: 0.05; 09: 0.07; 10: 0.07; 11: 0.089; 12: 0.0879; 13: 0.03; 14: 0.03 |  | 24
| 1 | 3 | 01: 0.08; 03: < 0.05; 04: 0.07; 05: 0.08; 06: 0.06; 07: 0.05; 09: 0.07; 10: 0.07; 11: 0.103; 12: 0.0771; 13: 0.02; 14: 0.04 |  | 16
| 1 | 4 | 01: 0.06; 03: < 0.05; 04: 0.07; 05: 0.08; 06: 0.06; 07: 0.11; 09: 0.06; 10: 0.07; 11: 0.091; 12: 0.0727; 13: 0.03; 14: 0.03 |  | 0.071317
| 1 | 5 | 01: 0.09; 03: < 0.05; 04: 0.07; 05: 0.07; 06: 0.07; 07: 0.08; 09: 0.07; 10: 0.07; 11: 0.089; 12: 0.0698; 13: 0.05; 14: 0.02 |  | 98
| 1 | 6 | 01: 0.06; 03: < 0.05; 04: 0.06; 05: 0.11; 06: 0.06; 07: 0.05; 09: 0.06; 10: 0.09; 11: 0.098; 12: 0.0649; 13: 0.03; 14: 0.02
| Mean |  | 01: 0.071667; 04: 0.066667; 05: 0.083333; 06: 0.061667; 07: 0.071667; 09: 0.068333; 10: 0.075; 11: 0.094667; 12: 0.0722; 13: 0.03; 14: 0.031667
| Median |  | 01: 0.07; 04: 0.07; 05: 0.08; 06: 0.06; 07: 0.065; 09: 0.07; 10: 0.07; 11: 0.0945; 12: 0.07125; 13: 0.03; 14: 0.03
| Std Dev. |  | 01: 0.01169; 04: 0.005164; 05: 0.015055; 06: 0.0040825; 07: 0.025626; 09: 0.0075277; 10: 0.0083666; 11: 0.0058195; 12: 0.0095892; 13: 0.010954; 14: 0.01169
| Rel.Std.Dev. |  | 01: 16.31%; 04: 7.746%; 05: 18.07%; 06: 6.62%; 07: 35.76%; 09: 11.02%; 10: 11.16%; 11: 6.147%; 12: 13.28%; 13: 36.51%; 14: 36.92%
| PDM3 |  | 01: 0.4908%; 04: -6.52%; 05: 16.85%; 06: -13.53%; 07: 0.4908%; 09: -4.183%; 10: 5.165%; 11: 32.74%; 12: 1.239%; 13: -57.93%; 14: -55.6%
| Z-Score (Absolute) |  | 01: 0; 03: NA; 04: 0.67; 05: 1.57; 06: 1.35; 07: 0; 09: 0.45; 10: 0.45; 11: 3.1; 12: 0.07; 13: 5.62; 14: 5.39
| Analytical results for Sb in OREAS 905 (Certified Value 1.96 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 26 | 30 | 32 | [8 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | ABL*MS | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 1.8356 | < 50 | 1.5 | 2 | 1.1 | 4 | 2.5 | < 50 | 2 | 1.9 | < 2 | 2.173 | 2.6 | 4.4 | 2.15 | [8 empty] | 1
| 1 | 2 | 1.888 | < 50 | 1.8 | 2 | 1 | 2 | 2.2 | < 50 | 1.4 | 1.3 | < 2 | 1.6 | 2.8 | 5.7 | 2.07 | [8 empty] | 0
| 1 | 3 | 2.4074 | < 50 | 1.4 | 2 | 0.8 | 2 | 2.2 | < 50 | 1.9 | 1.9 | < 2 | 1.576 | 2.1 | 5.1 | 2.12 | [8 empty] | 16
| 1 | 4 | 1.8021 | < 50 | 1.4 | 2 | 0.7 | 2 | 2.2 | < 50 | 2 | 2 | < 2 | 1.7 | 1.6 | 5.8 | 2.15 | [8 empty] | 1.9566
| 1 | 5 | 2.285 | < 50 | 1.3 | 2 | 0.9 | < 2 | 2.1 | < 50 | 1.8 | 1.8 | < 2 | 1.6 | 2.8 | 3 | 1.97 | [8 empty] | 99
| 1 | 6 | 1.7482 | < 50 | 1.1 | 2 | 0.8 | 2 | 2.3 | < 50 | 1.9 | 1.9 | < 2 | 1.9 | 2.2 | 3.6 | 2.2
| Mean |  | 01: 1.9944; 04: 1.4167; 05: 2; 06: 0.88333; 07: 2.4; 09: 2.25; 11: 1.8333; 12: 1.8; 14: 1.7582; 15: 2.35; 17: 4.6; 18: 2.11
| Median |  | 01: 1.8618; 04: 1.4; 05: 2; 06: 0.85; 07: 2; 09: 2.2; 11: 1.9; 12: 1.9; 14: 1.65; 15: 2.4; 17: 4.75; 18: 2.135
| Std Dev. |  | 01: 0.27899; 04: 0.23166; 05: 0; 06: 0.1472; 07: 0.89443; 09: 0.13784; 11: 0.22509; 12: 0.25298; 14: 0.23614; 15: 0.47223; 17: 1.1402; 18: 0.080747
| Rel.Std.Dev. |  | 01: 13.99%; 04: 16.35%; 05: 0%; 06: 16.66%; 07: 37.27%; 09: 6.126%; 11: 12.28%; 12: 14.05%; 14: 13.43%; 15: 20.09%; 17: 24.79%; 18: 3.827%
| PDM3 |  | 01: 1.929%; 04: -27.6%; 05: 2.217%; 06: -54.85%; 07: 22.66%; 09: 14.99%; 11: -6.301%; 12: -8.005%; 14: -10.14%; 15: 20.1%; 17: 135.1%; 18: 7.839%
| Z-Score (Absolute) |  | 0.01 | NA | 2.48 | 0.01 | 4.76 | 0.01 | 0.87 | NA | 0.33 | 0.42 | NA | 1.38 | 1.51 | 11.12 | 0.48
| Analytical results for Sc in OREAS 905 (Indicative Value 4.5 ppm) | [24 empty] | Indicative
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | [9 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [9 empty] | 2
2
| 1 | 1 | 7.7358 | < 5 | 5 | 5 | < 5 | < 10 | < 5 | < 5 | < 20 | 4.8 | 4 | 3.5 | 9 | < 5 | [9 empty] | 1
| 1 | 2 | 5.7629 | < 5 | < 5 | 5 | < 5 | < 10 | < 5 | < 5 | < 20 | 4.3 | 3 | 4.3 | 12 | < 5 | [9 empty] | 6
| 1 | 3 | 2.6771 | < 5 | 5 | 5 | < 5 | < 10 | < 5 | < 5 | < 20 | 4.3 | 4 | 3.4 | 12 | < 5 | [9 empty] | 16
| 1 | 4 | 7.2256 | < 5 | 7 | 5 | < 5 | < 10 | < 5 | < 5 | < 20 | 4.5 | 4 | 2.8 | 11 | < 5 | [9 empty] | 4.4997
| 1 | 5 | 5.9139 | < 5 | 6 | 5 | < 5 | < 10 | < 5 | < 5 | < 20 | 4.1 | 4 | 3.9 | 9 | < 5 | [9 empty] | 12
| 1 | 6 | 4.1753 | < 5 | < 5 | 5 | < 5 | < 10 | < 5 | < 5 | < 20 | 4.3 | 4 | 3.8 | 9 | < 5
| Mean |  | 01: 5.5818; 04: 5.75; 05: 5; 12: 4.3833; 13: 3.8333; 14: 3.6167; 15: 10.3333
| Median |  | 01: 5.8384; 04: 5.5; 05: 5; 12: 4.3; 13: 4; 14: 3.65; 15: 10
| Std Dev. |  | 01: 1.8921; 04: 0.95743; 05: 0; 12: 0.24014; 13: 0.40825; 14: 0.51153; 15: 1.5055
| Rel.Std.Dev. |  | 01: 33.9%; 04: 16.65%; 05: 0%; 12: 5.478%; 13: 10.65%; 14: 14.14%; 15: 14.57%
| PDM3 |  | 01: 24.05%; 04: 27.79%; 05: 11.12%; 12: -2.587%; 13: -14.81%; 14: -19.62%; 15: 129.6%
| Z-Score (Absolute) |  | 0.74 | NA | 0 | 0.27 | NA | NA | NA | NA | NA | 0.3 | 0.67 | 0.85 | 4.59 | NA
| Analytical results for Se in OREAS 905 (Indicative Value < 20 ppm) | [24 empty] | Indicative
Se | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 03; 03: 05; 04: 12; 05: 17; 06: 30 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: PF*MS; 05: PF*MS; 06: PF*MS |  | 1
1
| 1 | 1 | 01: < 20; 03: < 50; 04: 45.2; 05: < 20; 06: < 50 |  | 1
| 1 | 2 | 01: < 20; 03: < 50; 04: 47.6; 05: < 20; 06: < 50 |  | 7
| 1 | 3 | 01: < 20; 03: < 50; 04: 43.7; 05: < 20; 06: < 50 |  | 16
| 1 | 4 | 01: < 20; 03: < 50; 04: 45; 05: < 20; 06: < 50 |  | < 20
| 1 | 5 | 01: < 20; 03: < 50; 04: 30.6; 05: < 20; 06: < 50 |  | 13
| 1 | 6 | 01: < 20; 03: < 50; 04: 47.7; 05: < 20; 06: < 50
| Mean |  | 04: 43.3
| Median |  | 04: 45.1
| Std Dev. |  | 04: 6.4156
| Rel.Std.Dev. |  | 04: 14.82%
PDM3
Z-Score (Absolute) | bar: 01 to 06
| Analytical results for Si in OREAS 905 (Certified Value 32.33 wt.%) | [24 empty] | Certified
Si | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 06 | 07 | 09 | 11 | 12 | 13 | 17 | 20 | 22 | 26 | 30 | 31 | [9 empty] | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | BF*OES | PF*OES | BF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | [9 empty] | 2
3
| 1 | 1 | 33 | 30.9389 | 33.1 | 32.72 | 31.3 | 32.306 | 33.1 | 31.82 | 32.9 | 32.8813 | 31.211 | 31.8 | 24.2 | 33.89 | [9 empty] | 1
| 1 | 2 | 33.1 | 29.9673 | 32.58 | 33.52 | 31.7 | 32.723 | 33.3 | 31.92 | 33.1 | 33.2806 | 29.2273 | 31.8 | 23.6 | 34.41 | [9 empty] | 0
| 1 | 3 | 32.49 | 32.8629 | 32.49 | 32.82 | 30.8 | 32.554 | 32.9 | 31.65 | 32.4 | 33.3422 | 30.153 | 31.7 | 24.7 | 34.08 | [9 empty] | 16
| 1 | 4 | 32.49 | 31.6453 | 32.49 | 33.1 | 30.8 | 32.526 | 31.9 | 31.96 | 32.6 | 33.0167 | 30.682 | 31.7 | 25.5 | 35.01 | [9 empty] | 32.3295
| 1 | 5 | 32.72 | 30.1232 | 32.91 | 33.89 | 31.9 | 32.339 | 31.7 | 31.92 | 32.6 | 32.5403 | 30.4175 | 31.9 | 23.4 | 34.41 | [9 empty] | 100
| 1 | 6 | 32.72 | 30.5047 | 33 | 33.1 | 31 | 33.003 | 31.8 | 31.94 | 33.2 | 33.213 | 30.2853 | 31.9 | 22.6 | 34.92
| Mean |  | 32.7533 | 31.007 | 32.7617 | 33.1917 | 31.25 | 32.5752 | 32.45 | 31.8683 | 32.8 | 33.0457 | 30.3294 | 31.8 | 24 | 34.4533
| Median |  | 32.72 | 30.7218 | 32.745 | 33.1 | 31.15 | 32.54 | 32.4 | 31.92 | 32.75 | 33.1149 | 30.3514 | 31.8 | 23.9 | 34.41
| Std Dev. |  | 0.25375 | 1.0928 | 0.27345 | 0.44111 | 0.46797 | 0.25922 | 0.72595 | 0.11737 | 0.31623 | 0.30144 | 0.65681 | 0.089443 | 1.0257 | 0.44455
| Rel.Std.Dev. |  | 0.7747% | 3.524% | 0.8347% | 1.329% | 1.498% | 0.7958% | 2.237% | 0.3683% | 0.9641% | 0.9122% | 2.166% | 0.2813% | 4.274% | 1.29%
| PDM3 |  | 1.311% | -4.09% | 1.337% | 2.667% | -3.339% | 0.76% | 0.3729% | -1.426% | 1.455% | 2.215% | -6.187% | -1.638% | -25.76% | 6.569%
| Z-Score (Absolute) |  | 0.25 | 1.54 | 0.25 | 0.69 | 1.29 | 0.06 | 0.06 | 0.66 | 0.29 | 0.54 | 2.23 | 0.73 | 8.69 | 1.98
| Analytical results for Sm in OREAS 905 (Certified Value 7.64 ppm) | [24 empty] | Certified
Sm | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS |  | 2
3
| 1 | 1 | 01: 7.7; 03: 7.5; 04: 7.1; 05: 7.6; 06: 7.7; 07: 7.7; 09: 7.544; 10: 8.07; 11: 7.48 |  | 1
| 1 | 2 | 01: 7.6; 03: 8; 04: 7.1; 05: 7.4; 06: 7.8; 07: 7.1; 09: 7.768; 10: 7.9; 11: 7.65 |  | 0
| 1 | 3 | 01: 6.5; 03: 6.5; 04: 7.4; 05: 7.4; 06: 7.4; 07: 7.6; 09: 8.187; 10: 9.09; 11: 7.57 |  | 16
| 1 | 4 | 01: 8; 03: 8; 04: 7.8; 05: 7.3; 06: 7.8; 07: 8; 09: 8.278; 10: 8.68; 11: 7.52 |  | 7.6355
| 1 | 5 | 01: 7.3; 03: 7.5; 04: 7.6; 05: 6.5; 06: 7.6; 07: 7.8; 09: 7.948; 10: 8.76; 11: 7.57 |  | 101
| 1 | 6 | 01: 7; 03: 8; 04: 7.6; 05: 7.9; 06: 7.9; 07: 7.8; 09: 7.591; 10: 7.75; 11: 7.56
| Mean |  | 01: 7.35; 03: 7.5833; 04: 7.4333; 05: 7.35; 06: 7.7; 07: 7.6667; 09: 7.886; 10: 8.375; 11: 7.5583
| Median |  | 01: 7.45; 03: 7.75; 04: 7.5; 05: 7.4; 06: 7.75; 07: 7.75; 09: 7.858; 10: 8.375; 11: 7.565
| Std Dev. |  | 01: 0.53944; 03: 0.58452; 04: 0.28752; 05: 0.46797; 06: 0.17889; 07: 0.30768; 09: 0.30529; 10: 0.54069; 11: 0.057067
| Rel.Std.Dev. |  | 01: 7.339%; 03: 7.708%; 04: 3.868%; 05: 6.367%; 06: 2.323%; 07: 4.013%; 09: 3.871%; 10: 6.456%; 11: 0.755%
| PDM3 |  | 01: -3.74%; 03: -0.6838%; 04: -2.648%; 05: -3.74%; 06: 0.8442%; 07: 0.4076%; 09: 3.28%; 10: 9.684%; 11: -1.011%
| Z-Score (Absolute) |  | 01: 0.67; 03: 0.61; 04: 1.07; 05: 0.67; 06: 0.15; 07: 0; 09: 1.01; 10: 3.26; 11: 0.5
| Analytical results for Sn in OREAS 905 (Certified Value 7.49 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | 33 | [6 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | PF*OES | [6 empty] | 2
2
| 1 | 1 | < 100 | < 50 | < 100 | < 10 | 23 | < 10 | 5.8 | 6 | < 50 | 8 | < 100 | 8.8 | < 10 | 7.4 | 18.6 | 7.98 | < 50 | [6 empty] | 1
| 1 | 2 | < 100 | < 50 | < 100 | < 10 | 32 | < 10 | 5.4 | 7 | < 50 | 6 | < 100 | 8.8 | < 10 | 7.2 | 19.9 | 8.38 | 50 | [6 empty] | 0
| 1 | 3 | < 100 | < 50 | < 100 | < 10 | 13 | 10 | 4.9 | 6 | < 50 | 7 | < 100 | 7.9 | < 10 | 7 | 15.5 | 7.67 | < 50 | [6 empty] | 16
| 1 | 4 | < 100 | < 50 | < 100 | < 10 | 11 | < 10 | 5.1 | 6 | < 50 | 7 | < 100 | 7.7 | < 10 | 7.8 | 19.3 | 7.85 | 50 | [6 empty] | 7.4879
| 1 | 5 | < 100 | < 50 | < 100 | < 10 | 2 | < 10 | 5.8 | 6 | < 50 | 7 | < 100 | 8.7 | < 10 | 8.2 | 16 | 7.66 | < 50 | [6 empty] | 102
| 1 | 6 | < 100 | < 50 | < 100 | < 10 | 24 | < 10 | 5.5 | 6 | < 50 | 8 | < 100 | 7.8 | < 10 | 7.6 | 12.8 | 7.55 | 50
| Mean |  | 06: 17.5; 07: 10; 09: 5.4167; 10: 6.1667; 12: 7.1667; 14: 8.2833; 17: 7.5333; 18: 17.0167; 19: 7.8483; 22: 50
| Median |  | 06: 18; 07: 10; 09: 5.45; 10: 6; 12: 7; 14: 8.3; 17: 7.5; 18: 17.3; 19: 7.76; 22: 50
| Std Dev. |  | 06: 10.8213; 09: 0.3656; 10: 0.40825; 12: 0.75277; 14: 0.53448; 17: 0.43205; 18: 2.7257; 19: 0.30209; 22: 0
| Rel.Std.Dev. |  | 06: 61.84%; 09: 6.75%; 10: 6.62%; 12: 10.5%; 14: 6.452%; 17: 5.735%; 18: 16.02%; 19: 3.849%; 22: 0%
| PDM3 |  | 06: 133.7%; 07: 33.55%; 09: -27.66%; 10: -17.64%; 12: -4.289%; 14: 10.62%; 17: 0.6073%; 18: 127.3%; 19: 4.814%; 22: 567.7%
| Z-Score (Absolute) |  | NA | NA | NA | NA | 3.32 | 0.68 | 0.93 | 0.67 | NA | 0.32 | NA | 0.08 | NA | 0.19 | 3.15 | 0.08 | 14.75
| Analytical results for Sr in OREAS 905 (Certified Value 159 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | [7 empty] | ppm
|  |  | PF*MS | PF*OES | PF*OES | BF*MS | PF*OES/MS | PF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [7 empty] | 0
0
| 1 | 1 | 157.6567 | 159 | 155 | 175 | 170 | 156 | 153 | 153.9 | 183 | 160 | 160 | 160 | 125 | 171 | 190 | 166 | [7 empty] | 1
| 1 | 2 | 156.0931 | 158 | 165 | 175 | 180 | 157 | 155 | 156.1 | 326 | 160 | 164 | 161 | 110 | 167 | 200 | 162 | [7 empty] | 0
| 1 | 3 | 158.143 | 159 | 150 | 165 | 190 | 155 | 145 | 151.6 | 183 | 160 | 160 | 156 | 115 | 163 | 200 | 152 | [7 empty] | 16
| 1 | 4 | 151.6362 | 156 | 150 | 180 | 170 | 152 | 155 | 158.3 | 189 | 160 | 161 | 154 | 110 | 160 | 190 | 167 | [7 empty] | 158.6086
| 1 | 5 | 151.2937 | 160 | 145 | 195 | 180 | 157 | 134 | 154.5 | 258 | 160 | 158 | 156 | 110 | 164 | 180 | 158 | [7 empty] | 103
| 1 | 6 | 151.7946 | 158 | 155 | 190 | 160 | 156 | 154 | 152.4 | 277 | 160 | 159 | 168 | 105 | 159 | 190 | 151
| Mean |  | 154.4362 | 158.3333 | 153.3333 | 180 | 175 | 155.5 | 149.3333 | 154.4667 | 236 | 160 | 160.3333 | 159.1667 | 112.5 | 164 | 191.6667 | 159.3333
| Median |  | 153.9438 | 158.5 | 152.5 | 177.5 | 175 | 156 | 153.5 | 154.2 | 223.5 | 160 | 160 | 158 | 110 | 163.5 | 190 | 160
| Std Dev. |  | 3.2109 | 1.3663 | 6.8313 | 10.9545 | 10.4881 | 1.8708 | 8.4063 | 2.4566 | 60.1531 | 0 | 2.0656 | 5.0761 | 6.892 | 4.4721 | 7.5277 | 6.8605
| Rel.Std.Dev. |  | 2.079% | 0.8629% | 4.455% | 6.086% | 5.993% | 1.203% | 5.629% | 1.59% | 25.49% | 0% | 1.288% | 3.189% | 6.126% | 2.727% | 3.928% | 4.306%
| PDM3 |  | -2.631% | -0.1735% | -3.326% | 13.49% | 10.33% | -1.96% | -5.848% | -2.611% | 48.79% | 0.8773% | 1.087% | 0.3519% | -29.07% | 3.399% | 20.84% | 0.4569%
| Z-Score (Absolute) |  | 0.68 | 0.13 | 0.83 | 2.92 | 1.79 | 0.53 | 0.96 | 0.67 | 10.79 | 0.11 | 0.15 | 0.01 | 6.92 | 0.67 | 4.56 | 0.01
| Analytical results for Ta in OREAS 905 (Certified Value 1.38 ppm) | [24 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 05 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 26 | 30 | 32 | [8 empty] | ppm
|  |  | PF*MS | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 1.5184 | < 10 | < 2 | 2.6 | 2.2 | 1.5 | < 10 | 1.4 | 1.5 | < 5 | 1.227 | < 20 | 1.33 | 1.4 | 1.39 | [8 empty] | 1
| 1 | 2 | 1.4261 | < 10 | < 2 | 2.3 | 1.2 | 1.5 | < 10 | 1.2 | 1.5 | < 5 | 1.274 | < 20 | 1.47 | 1.5 | 1.36 | [8 empty] | 14
| 1 | 3 | 1.6278 | < 10 | < 2 | 1.9 | 1.3 | 1.5 | < 10 | 1.3 | 1.2 | < 5 | 1.331 | < 20 | 1.41 | 1.5 | 1.33 | [8 empty] | 16
| 1 | 4 | 1.5876 | < 10 | < 2 | 2.1 | 1.1 | 1.5 | < 10 | 1.5 | 1.2 | < 5 | 1.284 | < 20 | 1.37 | 1.5 | 1.29 | [8 empty] | 1.376
| 1 | 5 | 1.2746 | < 10 | < 2 | 2 | 0.9 | 1.5 | < 10 | 1.4 | 1.2 | < 5 | 1.545 | < 20 | 1.52 | 1.4 | 1.28 | [8 empty] | 104
| 1 | 6 | 1.5309 | < 10 | < 2 | 1.9 | 2.5 | 1.5 | < 10 | 1.3 | 1.2 | < 5 | 1.595 | < 20 | 1.5 | 1.5 | 1.38
| Mean |  | 01: 1.4942; 05: 2.1333; 06: 1.5333; 07: 1.5; 10: 1.35; 11: 1.3; 13: 1.376; 15: 1.4333; 17: 1.4667; 18: 1.3383
| Median |  | 01: 1.5247; 05: 2.05; 06: 1.25; 07: 1.5; 10: 1.35; 11: 1.2; 13: 1.3075; 15: 1.44; 17: 1.5; 18: 1.345
| Std Dev. |  | 01: 0.12758; 05: 0.27325; 06: 0.6532; 07: 0; 10: 0.10488; 11: 0.15492; 13: 0.15467; 15: 0.075542; 17: 0.05164; 18: 0.046224
| Rel.Std.Dev. |  | 01: 8.538%; 05: 12.81%; 06: 42.6%; 07: 0%; 10: 7.769%; 11: 11.92%; 13: 11.24%; 15: 5.27%; 17: 3.521%; 18: 3.454%
| PDM3 |  | 01: 8.596%; 05: 55.04%; 06: 11.44%; 07: 9.015%; 10: -1.886%; 11: -5.52%; 13: 0.00349%; 15: 4.17%; 17: 6.593%; 18: -2.734%
| Z-Score (Absolute) |  | 0.77 | NA | NA | 6.3 | 2.42 | 0.82 | NA | 0.47 | 0.91 | NA | 0.25 | NA | 0.25 | 0.54 | 0.57
| Analytical results for Tb in OREAS 905 (Certified Value 0.81 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS |  | 2
2
| 1 | 1 | 01: 0.8; 03: 1; 04: 0.92; 05: 0.7; 06: 0.7; 07: 0.83; 09: 0.777; 10: 0.84; 11: 0.889 |  | 1
| 1 | 2 | 01: 0.8; 03: 1; 04: 0.86; 05: 0.7; 06: 0.7; 07: 0.75; 09: 0.677; 10: 0.91; 11: 0.826 |  | 0
| 1 | 3 | 01: 0.7; 03: 1; 04: 0.84; 05: 0.7; 06: 0.7; 07: 0.82; 09: 0.708; 10: 0.84; 11: 0.862 |  | 16
| 1 | 4 | 01: 0.8; 03: 1; 04: 0.87; 05: 0.7; 06: 0.7; 07: 0.9; 09: 0.782; 10: 0.8; 11: 0.814 |  | 0.81381
| 1 | 5 | 01: 0.8; 03: 1; 04: 0.9; 05: 0.6; 06: 0.7; 07: 0.84; 09: 0.678; 10: 0.92; 11: 0.844 |  | 105
| 1 | 6 | 01: 0.8; 03: 1; 04: 0.85; 05: 0.7; 06: 0.8; 07: 0.87; 09: 0.713; 10: 0.86; 11: 0.856
| Mean |  | 01: 0.78333; 03: 1; 04: 0.87333; 05: 0.68333; 06: 0.71667; 07: 0.835; 09: 0.7225; 10: 0.86167; 11: 0.8485
| Median |  | 01: 0.8; 03: 1; 04: 0.865; 05: 0.7; 06: 0.7; 07: 0.835; 09: 0.7105; 10: 0.85; 11: 0.85
| Std Dev. |  | 01: 0.040825; 03: 0; 04: 0.030768; 05: 0.040825; 06: 0.040825; 07: 0.050892; 09: 0.046608; 10: 0.045789; 11: 0.026816
| Rel.Std.Dev. |  | 01: 5.212%; 03: 0%; 04: 3.523%; 05: 5.974%; 06: 5.696%; 07: 6.095%; 09: 6.451%; 10: 5.314%; 11: 3.16%
| PDM3 |  | 01: -3.746%; 03: 22.88%; 04: 7.314%; 05: -16.03%; 06: -11.94%; 07: 2.603%; 09: -11.22%; 10: 5.88%; 11: 4.262%
| Z-Score (Absolute) |  | 01: 0.67; 03: 2.15; 04: 0.5; 05: 1.98; 06: 1.54; 07: 0; 09: 1.47; 10: 0.35; 11: 0.18
| Analytical results for Te in OREAS 905 (Indicative Value < 1 ppm) | [24 empty] | Indicative
Te | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 12; 06: 26; 07: 30 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*MS; 06: ABL*MS; 07: PF*MS |  | 2
2
| 1 | 1 | 01: < 1; 03: < 1; 04: < 5; 05: < 6; 06: < 0.2; 07: 7 |  | 1
| 1 | 2 | 01: < 1; 03: < 1; 04: < 5; 05: < 6; 06: < 0.2; 07: 8 |  | 8
| 1 | 3 | 01: < 1; 03: < 1; 04: < 5; 05: < 6; 06: < 0.2; 07: 8 |  | 16
| 1 | 4 | 01: < 1; 03: < 1; 04: < 5; 05: < 6; 06: < 0.2; 07: 9 |  | < 1
| 1 | 5 | 01: < 1; 03: < 1; 04: < 5; 05: < 6; 06: < 0.2; 07: 6 |  | 14
| 1 | 6 | 01: < 1; 03: < 1; 04: < 5; 05: < 6; 06: < 0.2; 07: 7
| Mean |  | 07: 7.5
| Median |  | 07: 7.5
| Std Dev. |  | 07: 1.0488
| Rel.Std.Dev. |  | 07: 13.98%
PDM3
Z-Score (Absolute) | bar: 01 to 07
| Analytical results for Th in OREAS 905 (Certified Value 14.7 ppm) | [24 empty] | Certified
Th | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 09: 15; 10: 17; 11: 20; 12: 22; 13: 26; 14: 30; 15: 32 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: BF*MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*MS; 12: PF*OES/MS; 13: ABL*MS; 14: PF*MS; 15: PF*OES/MS |  | 1
2
| 1 | 1 | 01: 14.7223; 03: 14.4; 04: 14.5; 05: 7.8; 06: 14; 07: 14.2; 09: 14.9; 10: 14.7; 11: 15.3439; 12: < 20; 13: 14.9; 14: 25.7; 15: 15.7 |  | 1
| 1 | 2 | 01: 14.3755; 03: 14.7; 04: 15.5; 05: 9.5; 06: 13.9; 07: 14.1; 09: 13; 10: 14.8; 11: 14.9234; 12: < 20; 13: 16.5; 14: 24.7; 15: 15.3 |  | 0
| 1 | 3 | 01: 14.2619; 03: 13.4; 04: 14.5; 05: 11; 06: 13.4; 07: 14.1; 09: 14.7; 10: 14.2; 11: 14.7374; 12: < 20; 13: 15.4; 14: 25; 15: 15.1 |  | 16
| 1 | 4 | 01: 14.7426; 03: 15.4; 04: 16; 05: 10; 06: 13.2; 07: 14.5; 09: 15; 10: 14.6; 11: 14.8084; 12: < 20; 13: 14.9; 14: 24.1; 15: 15 |  | 14.6884
| 1 | 5 | 01: 14.7843; 03: 14; 04: 15.5; 05: 11; 06: 12.1; 07: 14.3; 09: 14.8; 10: 14.5; 11: 14.4598; 12: < 20; 13: 16.2; 14: 24.5; 15: 15.2 |  | 106
| 1 | 6 | 01: 14.3574; 03: 14; 04: 15.5; 05: 11.3; 06: 14; 07: 14.2; 09: 15; 10: 14.5; 11: 14.6084; 12: < 20; 13: 14.7; 14: 24.3; 15: 14.7
| Mean |  | 01: 14.5407; 03: 14.3167; 04: 15.25; 05: 10.1; 06: 13.4333; 07: 14.2333; 09: 14.5667; 10: 14.55; 11: 14.8136; 13: 15.4333; 14: 24.7167; 15: 15.1667
| Median |  | 01: 14.5489; 03: 14.2; 04: 15.5; 05: 10.5; 06: 13.65; 07: 14.2; 09: 14.85; 10: 14.55; 11: 14.7729; 13: 15.15; 14: 24.6; 15: 15.15
| Std Dev. |  | 01: 0.23309; 03: 0.68823; 04: 0.61237; 05: 1.3206; 06: 0.73394; 07: 0.15055; 09: 0.77632; 10: 0.20736; 11: 0.30542; 13: 0.75277; 14: 0.57417; 15: 0.33267
| Rel.Std.Dev. |  | 01: 1.603%; 03: 4.807%; 04: 4.016%; 05: 13.08%; 06: 5.464%; 07: 1.058%; 09: 5.329%; 10: 1.425%; 11: 2.062%; 13: 4.878%; 14: 2.323%; 15: 2.193%
| PDM3 |  | 01: -1.006%; 03: -2.531%; 04: 3.823%; 05: -31.24%; 06: -8.545%; 07: -3.098%; 09: -0.8289%; 10: -0.9424%; 11: 0.8519%; 13: 5.071%; 14: 68.27%; 15: 3.256%
| Z-Score (Absolute) |  | 01: 0.2; 03: 0.53; 04: 0.82; 05: 6.62; 06: 1.42; 07: 0.65; 09: 0.29; 10: 0.19; 11: 0.19; 12: NA; 13: 1.09; 14: 14.5; 15: 0.7
| Analytical results for Ti in OREAS 905 (Certified Value 0.128 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 26 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*OES | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES |  | 3
3
| 1 | 1 | 0.132 | 0.11738 | < 0.01 | 0.14 | 0.132 | 0.12 | 0.125 | 0.12 | 0.126 | 0.14 | 0.13 | 0.1 | 0.12 | 0.13 | 0.135 | 0.12858 | 0.0946 | 0.13 | 0.13 | 0.126 | 0.122 | 0.12 |  | 1
| 1 | 2 | 0.132 | 0.11821 | < 0.01 | 0.13 | 0.132 | 0.132 | 0.125 | 0.12 | 0.126 | 0.14 | 0.13 | 0.12 | 0.12 | 0.13 | 0.135 | 0.12837 | 0.0848 | 0.13 | 0.13 | 0.126 | 0.119 | 0.12 |  | 0
| 1 | 3 | 0.132 | 0.12325 | < 0.01 | 0.13 | 0.132 | 0.132 | 0.125 | 0.13 | 0.126 | 0.13 | 0.13 | 0.09 | 0.12 | 0.13 | 0.129 | 0.12964 | 0.0889 | 0.13 | 0.13 | 0.126 | 0.121 | 0.125 |  | 16
| 1 | 4 | 0.132 | 0.12073 | < 0.01 | 0.13 | 0.132 | 0.132 | 0.13 | 0.12 | 0.126 | 0.13 | 0.13 | 0.1 | 0.12 | 0.13 | 0.128 | 0.13032 | 0.0957 | 0.13 | 0.13 | 0.132 | 0.123 | 0.122 |  | 0.12759
| 1 | 5 | 0.132 | 0.11099 | < 0.01 | 0.13 | 0.132 | 0.132 | 0.135 | 0.12 | 0.126 | 0.13 | 0.13 | 0.1 | 0.12 | 0.13 | 0.131 | 0.12808 | 0.0902 | 0.13 | 0.12 | 0.132 | 0.12 | 0.133 |  | 107
| 1 | 6 | 0.132 | 0.11786 | < 0.01 | 0.13 | 0.132 | 0.132 | 0.13 | 0.12 | 0.126 | 0.13 | 0.13 | 0.11 | 0.12 | 0.13 | 0.129 | 0.12917 | 0.0914 | 0.13 | 0.12 | 0.126 | 0.12 | 0.127
| Mean |  | 0.132 | 0.11807 |  | 0.13167 | 0.132 | 0.13 | 0.12833 | 0.12167 | 0.126 | 0.13333 | 0.13 | 0.10333 | 0.12 | 0.13 | 0.13117 | 0.12903 | 0.090933 | 0.13 | 0.12667 | 0.128 | 0.12083 | 0.1245
| Median |  | 0.132 | 0.11803 |  | 0.13 | 0.132 | 0.132 | 0.1275 | 0.12 | 0.126 | 0.13 | 0.13 | 0.1 | 0.12 | 0.13 | 0.13 | 0.12887 | 0.0908 | 0.13 | 0.13 | 0.126 | 0.1205 | 0.1235
| Std Dev. |  | 0 | 0.0041138 |  | 0.0040825 | 0 | 0.004899 | 0.0040825 | 0.0040825 | 0 | 0.005164 | 0 | 0.010328 | 0 | 0 | 0.0031252 | 0.00084873 | 0.0039667 | 0 | 0.005164 | 0.0030984 | 0.001472 | 0.00501
| Rel.Std.Dev. |  | 0% | 3.484% |  | 3.101% | 0% | 3.768% | 3.181% | 3.355% | 0% | 3.873% | 0% | 9.995% | 0% | 0% | 2.383% | 0.6578% | 4.362% | 0% | 4.077% | 2.421% | 1.218% | 4.024%
| PDM3 |  | 3.46% | -7.459% |  | 3.199% | 3.46% | 1.893% | 0.5862% | -4.639% | -1.243% | 4.505% | 1.893% | -19.01% | -5.945% | 1.893% | 2.807% | 1.13% | -28.73% | 1.893% | -0.7201% | 0.325% | -5.292% | -2.418%
| Z-Score (Absolute) |  | 0.72 | 1.79 | NA | 0.67 | 0.72 | 0.31 | 0 | 1.35 | 0.49 | 1.01 | 0.34 | 5.06 | 1.69 | 0.34 | 0.57 | 0.14 | 7.57 | 0.34 | 0.34 | 0.09 | 1.52 | 0.78
| Analytical results for Tl in OREAS 905 (Certified Value 0.74 ppm) | [24 empty] | Certified
Tl | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 15; 09: 20; 10: 26; 11: 30; 12: 32 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*MS; 12: PF*OES/MS |  | 2
3
| 1 | 1 | 01: 0.75203; 03: 0.7; 04: < 3; 05: 0.5; 06: 0.7; 07: 0.7; 09: 0.735; 10: 1; 11: 0.4; 12: 0.76 |  | 1
| 1 | 2 | 01: 0.72561; 03: 0.8; 04: < 3; 05: 0.5; 06: 0.7; 07: 0.8; 09: 0.81; 10: 0.8; 11: 0.4; 12: 0.73 |  | 0
| 1 | 3 | 01: 0.73024; 03: 0.8; 04: < 3; 05: 0.6; 06: 0.7; 07: 0.7; 09: 0.813; 10: 1; 11: 0.4; 12: 0.74 |  | 16
| 1 | 4 | 01: 0.80968; 03: 0.8; 04: < 3; 05: 0.7; 06: 0.7; 07: 0.8; 09: 0.698; 10: 0.8; 11: 0.4; 12: 0.74 |  | 0.7437
| 1 | 5 | 01: 0.75929; 03: 0.7; 04: < 3; 05: 0.6; 06: 0.7; 07: 0.7; 09: 0.781; 10: 1; 11: 0.4; 12: 0.7 |  | 108
| 1 | 6 | 01: 0.70343; 03: 0.8; 04: < 3; 05: 0.6; 06: 0.7; 07: 0.8; 09: 0.716; 10: 1; 11: 0.4; 12: 0.77
| Mean |  | 01: 0.74671; 03: 0.76667; 05: 0.58333; 06: 0.7; 07: 0.75; 09: 0.75883; 10: 0.93333; 11: 0.4; 12: 0.74
| Median |  | 01: 0.74113; 03: 0.8; 05: 0.6; 06: 0.7; 07: 0.75; 09: 0.758; 10: 1; 11: 0.4; 12: 0.74
| Std Dev. |  | 01: 0.036719; 03: 0.05164; 05: 0.075277; 06: 1.2162e-16; 07: 0.054772; 09: 0.049289; 10: 0.10328; 11: 6.0809e-17; 12: 0.024495
| Rel.Std.Dev. |  | 01: 4.917%; 03: 6.736%; 05: 12.9%; 06: 1.737e-14%; 07: 7.303%; 09: 6.495%; 10: 11.07%; 11: 1.52e-14%; 12: 3.31%
| PDM3 |  | 01: 0.4048%; 03: 3.088%; 05: -21.56%; 06: -5.876%; 07: 0.8468%; 09: 2.035%; 10: 25.5%; 11: -46.22%; 12: -0.4978%
| Z-Score (Absolute) |  | 01: 0; 03: 0.67; 04: NA; 05: 5.52; 06: 1.58; 07: 0.11; 09: 0.41; 10: 6.31; 11: 11.72; 12: 0.23
| Analytical results for Tm in OREAS 905 (Certified Value 0.14 ppm) | [24 empty] | Certified
Tm | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 32 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*OES/MS |  | 2
2
| 1 | 1 | 01: < 0.25; 03: < 1; 04: < 0.05; 05: 0.1; 06: 0.2; 07: 0.14; 09: 0.15; 10: 0.16; 11: 0.144 |  | 1
| 1 | 2 | 01: < 0.25; 03: < 1; 04: < 0.05; 05: 0.1; 06: 0.1; 07: 0.13; 09: 0.113; 10: 0.14; 11: 0.148 |  | 0
| 1 | 3 | 01: < 0.25; 03: < 1; 04: < 0.05; 05: 0.1; 06: 0.2; 07: 0.14; 09: 0.127; 10: 0.17; 11: 0.146 |  | 16
| 1 | 4 | 01: < 0.25; 03: < 1; 04: < 0.05; 05: 0.1; 06: 0.1; 07: 0.13; 09: 0.136; 10: 0.13; 11: 0.145 |  | 0.14077
| 1 | 5 | 01: < 0.25; 03: < 1; 04: < 0.05; 05: 0.1; 06: 0.1; 07: 0.14; 09: 0.131; 10: 0.13; 11: 0.145 |  | 109
| 1 | 6 | 01: < 0.25; 03: < 1; 04: < 0.05; 05: 0.1; 06: 0.2; 07: 0.14; 09: 0.124; 10: 0.12; 11: 0.144
| Mean |  | 05: 0.1; 06: 0.15; 07: 0.13667; 09: 0.13017; 10: 0.14167; 11: 0.14533
| Median |  | 05: 0.1; 06: 0.15; 07: 0.14; 09: 0.129; 10: 0.135; 11: 0.145
| Std Dev. |  | 05: 1.5202e-17; 06: 0.054772; 07: 0.005164; 09: 0.012416; 10: 0.019408; 11: 0.0015055
| Rel.Std.Dev. |  | 05: 1.52e-14%; 06: 36.51%; 07: 3.779%; 09: 9.539%; 10: 13.7%; 11: 1.036%
| PDM3 |  | 05: -28.96%; 06: 6.559%; 07: -2.913%; 09: -7.53%; 10: 0.6394%; 11: 3.244%
| Z-Score (Absolute) |  | 01: NA; 03: NA; 04: NA; 05: 3.48; 06: 0.96; 07: 0.22; 09: 0.8; 10: 0.22; 11: 0.55
| Analytical results for U in OREAS 905 (Certified Value 5 ppm) | [24 empty] | Certified
U | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 10; 06: 12; 07: 13; 09: 15; 10: 20; 11: 22; 12: 26; 13: 30; 14: 32 |  | ppm
|  |  | 01: PF*MS; 03: PF*MS; 04: BF*MS; 05: PF*OES/MS; 06: PF*MS; 07: BF*MS; 09: PF*OES/MS; 10: PF*MS; 11: PF*OES/MS; 12: ABL*MS; 13: PF*MS; 14: PF*OES/MS |  | 2
3
| 1 | 1 | 01: 5.1705; 03: 5.1; 04: 4.5; 05: 4.07; 06: 5; 07: 4.8; 09: 5.43; 10: 5.1184; 11: < 50; 12: 5.29; 13: 6.8; 14: 5.52 |  | 1
| 1 | 2 | 01: 5.071; 03: 5.2; 04: 4.5; 05: 3.97; 06: 5; 07: 5; 09: 4.84; 10: 5.0096; 11: < 50; 12: 5.19; 13: 6.5; 14: 5.38 |  | 0
| 1 | 3 | 01: 5.4219; 03: 4.8; 04: 4; 05: 4.26; 06: 4.8; 07: 4.7; 09: 5.31; 10: 5.0784; 11: < 50; 12: 5.52; 13: 7.3; 14: 5.46 |  | 16
| 1 | 4 | 01: 5.5938; 03: 5.2; 04: 5; 05: 6.08; 06: 4.9; 07: 4.9; 09: 5.51; 10: 4.9256; 11: < 50; 12: 5.3; 13: 6.7; 14: 5.31 |  | 5.0041
| 1 | 5 | 01: 5.3412; 03: 5; 04: 6.5; 05: 4.06; 06: 4.6; 07: 4.9; 09: 5.36; 10: 5.0008; 11: < 50; 12: 5.54; 13: 6.8; 14: 5.43 |  | 110
| 1 | 6 | 01: 4.8991; 03: 4.9; 04: 5; 05: 4.38; 06: 5.5; 07: 4.9; 09: 5.48; 10: 4.9584; 11: < 50; 12: 5.09; 13: 6.9; 14: 5.43
| Mean |  | 01: 5.2496; 03: 5.0333; 04: 4.9167; 05: 4.47; 06: 4.9667; 07: 4.8667; 09: 5.3217; 10: 5.0152; 12: 5.3217; 13: 6.8333; 14: 5.4217
| Median |  | 01: 5.2558; 03: 5.05; 04: 4.75; 05: 4.165; 06: 4.95; 07: 4.9; 09: 5.395; 10: 5.0052; 12: 5.295; 13: 6.8; 14: 5.43
| Std Dev. |  | 01: 0.25202; 03: 0.1633; 04: 0.8612; 05: 0.8028; 06: 0.30111; 07: 0.10328; 09: 0.24734; 10: 0.072302; 12: 0.1786; 13: 0.26583; 14: 0.071391
| Rel.Std.Dev. |  | 01: 4.801%; 03: 3.244%; 04: 17.52%; 05: 17.96%; 06: 6.063%; 07: 2.122%; 09: 4.648%; 10: 1.442%; 12: 3.356%; 13: 3.89%; 14: 1.317%
| PDM3 |  | 01: 4.906%; 03: 0.5847%; 04: -1.747%; 05: -10.67%; 06: -0.7476%; 07: -2.746%; 09: 6.347%; 10: 0.2223%; 12: 6.347%; 13: 36.56%; 14: 8.345%
| Z-Score (Absolute) |  | 01: 0.51; 03: 0; 04: 1.01; 05: 2.07; 06: 0.16; 07: 0.39; 09: 0.9; 10: 0.04; 11: NA; 12: 0.67; 13: 4.21; 14: 0.91
| Analytical results for V in OREAS 905 (Certified Value 11.3 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | [7 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | BF*OES | PF*OES/MS | BF*OES | PF*MS | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | ABL*MS | PF*OES | PF*OES/MS | [7 empty] | 1
1
| 1 | 1 | < 20 | < 10 | < 10 | < 20 | 10 | 14 | 12 | 17 | 12 | 12 | < 20 | 11.6 | < 20 | 9.8 | < 10 | 10.5 | [7 empty] | 1
| 1 | 2 | < 20 | < 10 | < 10 | < 20 | < 10 | 12 | 11 | 11 | 14 | 10 | < 20 | 11.1 | < 20 | 10.1 | < 10 | 10.2 | [7 empty] | 0
| 1 | 3 | < 20 | < 10 | < 10 | < 20 | 10 | 14 | 10 | 17 | 11 | 12 | < 20 | 10.7 | < 20 | 9.1 | < 10 | 10.3 | [7 empty] | 16
| 1 | 4 | < 20 | < 10 | < 10 | < 20 | 10 | 16 | 12 | 12 | 12 | 12 | < 20 | 11.1 | < 20 | 9.3 | < 10 | 11.2 | [7 empty] | 11.2958
| 1 | 5 | < 20 | < 10 | < 10 | < 20 | 10 | 15 | 10 | 18 | 12 | 11 | < 20 | 10.5 | < 20 | 10.8 | < 10 | 12.1 | [7 empty] | 111
| 1 | 6 | < 20 | < 10 | < 10 | < 20 | < 10 | 12 | 8 | 18 | 14 | 12 | < 20 | 10.8 | < 20 | 11.2 | < 10 | 11.8
| Mean |  | 06: 10; 07: 13.8333; 09: 10.5; 10: 15.5; 11: 12.5; 12: 11.5; 14: 10.9667; 17: 10.05; 19: 11.0167
| Median |  | 06: 10; 07: 14; 09: 10.5; 10: 17; 11: 12; 12: 12; 14: 10.95; 17: 9.95; 19: 10.85
| Std Dev. |  | 06: 0; 07: 1.6021; 09: 1.5166; 10: 3.1464; 11: 1.2247; 12: 0.83666; 14: 0.38816; 17: 0.82644; 19: 0.8085
| Rel.Std.Dev. |  | 06: 0%; 07: 11.58%; 09: 14.44%; 10: 20.3%; 11: 9.798%; 12: 7.275%; 14: 3.539%; 17: 8.223%; 19: 7.339%
| PDM3 |  | 06: -11.47%; 07: 22.46%; 09: -7.045%; 10: 37.22%; 11: 10.66%; 12: 1.807%; 14: -2.914%; 17: -11.03%; 19: -2.471%
| Z-Score (Absolute) |  | NA | NA | NA | NA | 0.71 | 1.96 | 0.36 | 3.13 | 1.03 | 0.34 | NA | 0.03 | NA | 0.67 | NA | 0
| Analytical results for W in OREAS 905 (Certified Value 3.02 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | [8 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | [8 empty] | 2
2
| 1 | 1 | 2.8549 | < 50 | 3 | 3 | 9 | 1.1 | 4 | < 50 | 3 | 3 | < 5 | < 20 | 2.45 | 3.3 | 3.06 | [8 empty] | 1
| 1 | 2 | 2.8178 | < 50 | < 2 | 3 | 8 | 4.2 | 4 | < 50 | 3 | 3 | < 5 | < 20 | 2.7 | 3.2 | 3.06 | [8 empty] | 0
| 1 | 3 | 3.2897 | < 50 | 3 | 3 | 8 | 1.2 | 3 | < 50 | 3 | 2 | < 5 | < 20 | 2.9 | 3.3 | 2.79 | [8 empty] | 16
| 1 | 4 | 2.8372 | < 50 | 3 | 6 | 7 | 2.7 | 3 | < 50 | 3 | 4 | < 5 | < 20 | 2.4 | 3.4 | 2.77 | [8 empty] | 3.0182
| 1 | 5 | 3.1318 | < 50 | 2 | 6 | 6 | 2.2 | 3 | < 50 | 3 | 3 | < 5 | < 20 | 2.85 | 2.9 | 2.85 | [8 empty] | 112
| 1 | 6 | 3.0521 | < 50 | 3 | 6 | 6 | < 0.7 | 3 | < 50 | 3 | 4 | < 5 | < 20 | 2.5 | 3.3 | 2.87
| Mean |  | 01: 2.9972; 04: 2.8; 05: 4.5; 06: 7.3333; 07: 2.28; 09: 3.3333; 11: 3; 12: 3.1667; 15: 2.6333; 17: 3.2333; 18: 2.9
| Median |  | 01: 2.9535; 04: 3; 05: 4.5; 06: 7.5; 07: 2.2; 09: 3; 11: 3; 12: 3; 15: 2.6; 17: 3.3; 18: 2.86
| Std Dev. |  | 01: 0.19221; 04: 0.44721; 05: 1.6432; 06: 1.2111; 07: 1.2677; 09: 0.5164; 11: 0; 12: 0.75277; 15: 0.2137; 17: 0.17512; 18: 0.12931
| Rel.Std.Dev. |  | 01: 6.413%; 04: 15.97%; 05: 36.51%; 06: 16.51%; 07: 55.6%; 09: 15.49%; 11: 0%; 12: 23.77%; 15: 8.115%; 17: 5.416%; 18: 4.459%
| PDM3 |  | 01: -0.6949%; 04: -7.23%; 05: 49.09%; 06: 143%; 07: -24.46%; 09: 10.44%; 11: -0.6034%; 12: 4.919%; 15: -12.75%; 17: 7.127%; 18: -3.917%
| Z-Score (Absolute) |  | 0.1 | NA | 0.67 | 4.24 | 12.43 | 0 | 0.87 | NA | 0.1 | 0.39 | NA | NA | 1.16 | 0.77 | 0.39
| Analytical results for Y in OREAS 905 (Certified Value 16.6 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 03 | 04 | 05 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 20 | 22 | 26 | 30 | 32 | [6 empty] | ppm
|  |  | PF*MS | PF*OES | PF*MS | BF*MS | PF*OES/MS | BF*OES | PF*MS | BF*MS | PF*OES | PF*OES/MS | PF*MS | PF*OES | PF*MS | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | [6 empty] | 1
2
| 1 | 1 | 14.4871 | 16 | 16.5 | 17 | 18.4 | 16.1 | 17 | 16.5 | 16 | 17.2 | 15.4 | 18.7 | 18 | 14 | 18.4 | 11.5 | 17.3 | [6 empty] | 1
| 1 | 2 | 14.6773 | 15 | 15.9 | 17 | 17.9 | 16.1 | 17.2 | 16.7 | 18 | 15.6 | 16.2 | 18.3 | 16.633 | 13 | 18.9 | 11.5 | 16.9 | [6 empty] | 0
| 1 | 3 | 13.7083 | 15 | 15.8 | 16 | 18.1 | 16 | 16 | 16.5 | 15 | 16.7 | 15.4 | 17.7 | 17.367 | 14 | 20.4 | 11.7 | 17.3 | [6 empty] | 16
| 1 | 4 | 14.8104 | 15 | 16.7 | 18 | 17.7 | 16.3 | 16.8 | 16.7 | 18 | 17.3 | 15.4 | 17.5 | 16.838 | 15 | 18.5 | 10.9 | 17.1 | [6 empty] | 16.5676
| 1 | 5 | 14.5998 | 14 | 16.5 | 17 | 17.6 | 16 | 15.2 | 16.5 | 16 | 17.1 | 15.8 | 17.7 | 17.008 | 14 | 19.4 | 11.1 | 17.1 | [6 empty] | 113
| 1 | 6 | 14.0733 | 16 | 15.6 | 19 | 17.8 | 16.1 | 17.3 | 16.9 | 19 | 17.3 | 15.1 | 16.6 | 16.648 | 15 | 19 | 11.1 | 17.2
| Mean |  | 14.3927 | 15.1667 | 16.1667 | 17.3333 | 17.9167 | 16.1 | 16.5833 | 16.6333 | 17 | 16.8667 | 15.55 | 17.75 | 17.0823 | 14.1667 | 19.1 | 11.3 | 17.15
| Median |  | 14.5435 | 15 | 16.2 | 17 | 17.85 | 16.1 | 16.9 | 16.6 | 17 | 17.15 | 15.4 | 17.7 | 16.923 | 14 | 18.95 | 11.3 | 17.15
| Std Dev. |  | 0.41898 | 0.75277 | 0.45461 | 1.0328 | 0.29269 | 0.10954 | 0.82077 | 0.1633 | 1.5492 | 0.65929 | 0.38859 | 0.72042 | 0.52515 | 0.75277 | 0.73212 | 0.30984 | 0.15166
| Rel.Std.Dev. |  | 2.911% | 4.963% | 2.812% | 5.958% | 1.634% | 0.6804% | 4.949% | 0.9818% | 9.113% | 3.909% | 2.499% | 4.059% | 3.074% | 5.314% | 3.833% | 2.742% | 0.8843%
| PDM3 |  | -13.13% | -8.456% | -2.42% | 4.622% | 8.143% | -2.822% | 0.09492% | 0.3967% | 2.61% | 1.805% | -6.142% | 7.137% | 3.107% | -14.49% | 15.29% | -31.79% | 3.515%
| Z-Score (Absolute) |  | 2.16 | 1.41 | 0.45 | 0.67 | 1.24 | 0.51 | 0.05 | 0 | 0.35 | 0.47 | 1.17 | 1.08 | 0.43 | 2.38 | 2.38 | 5.14 | 0.5
| Analytical results for Yb in OREAS 905 (Certified Value 0.76 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab |  | 1
| No | No | 01: 05; 03: 09; 04: 10; 05: 12; 06: 13; 07: 15; 09: 20; 10: 26; 11: 30; 12: 32 |  | ppm
|  |  | 01: PF*MS; 03: BF*MS; 04: PF*OES/MS; 05: PF*MS; 06: BF*MS; 07: PF*OES/MS; 09: PF*MS; 10: ABL*MS; 11: PF*MS; 12: PF*OES/MS |  | 2
3
| 1 | 1 | 01: 0.7; 03: < 1; 04: 0.8; 05: 0.7; 06: 0.7; 07: 0.8; 09: 0.717; 10: 0.81; 11: 1.7; 12: 0.79 |  | 1
| 1 | 2 | 01: < 0.5; 03: < 1; 04: 0.8; 05: 0.7; 06: 0.8; 07: 0.7; 09: 0.801; 10: 0.91; 11: 1.7; 12: 0.746 |  | 0
| 1 | 3 | 01: 0.7; 03: < 1; 04: 0.9; 05: 0.7; 06: 0.7; 07: 0.7; 09: 0.724; 10: 0.82; 11: 1.7; 12: 0.792 |  | 16
| 1 | 4 | 01: 0.5; 03: < 1; 04: 0.9; 05: 0.7; 06: 0.7; 07: 0.8; 09: 0.782; 10: 0.85; 11: 1.7; 12: 0.757 |  | 0.75552
| 1 | 5 | 01: 0.6; 03: < 1; 04: 0.8; 05: 0.6; 06: 0.8; 07: 0.8; 09: 0.689; 10: 0.84; 11: 1.6; 12: 0.774 |  | 114
| 1 | 6 | 01: 0.6; 03: < 1; 04: 0.9; 05: 0.7; 06: 0.8; 07: 0.8; 09: 0.753; 10: 0.83; 11: 1.8; 12: 0.76
| Mean |  | 01: 0.62; 04: 0.85; 05: 0.68333; 06: 0.75; 07: 0.76667; 09: 0.74433; 10: 0.84333; 11: 1.7; 12: 0.76983
| Median |  | 01: 0.6; 04: 0.85; 05: 0.7; 06: 0.75; 07: 0.8; 09: 0.7385; 10: 0.835; 11: 1.7; 12: 0.767
| Std Dev. |  | 01: 0.083666; 04: 0.054772; 05: 0.040825; 06: 0.054772; 07: 0.05164; 09: 0.042255; 10: 0.03559; 11: 0.063246; 12: 0.018681
| Rel.Std.Dev. |  | 01: 13.49%; 04: 6.444%; 05: 5.974%; 06: 7.303%; 07: 6.736%; 09: 5.677%; 10: 4.22%; 11: 3.72%; 12: 2.427%
| PDM3 |  | 01: -17.94%; 04: 12.51%; 05: -9.555%; 06: -0.7307%; 07: 1.475%; 09: -1.481%; 10: 11.62%; 11: 125%; 12: 1.894%
| Z-Score (Absolute) |  | 01: 1.24; 03: NA; 04: 0.89; 05: 0.89; 06: 0.18; 07: 0; 09: 0.24; 10: 0.67; 11: 9.94; 12: 0.03
| Analytical results for Zn in OREAS 905 (Certified Value 139 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 03 | 04 | 05 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 26 | 30 | 32 | 33 | [7 empty] | ppm
|  |  | PF*OES | PF*OES | PF*OES | PF*OES/MS | PF*OES | PF*MS | BF*OES | PF*OES | PF*OES/MS | PF*OES | PF*OES | PF*OES/MS | ABL*MS | PF*MS | PF*OES/MS | PF*OES | [7 empty] | 0
0
| 1 | 1 | 136.7035 | 134 | 145 | 160 | 131 | 120 | 157 | 130 | 132 | 135 | 141 | 103 | 135 | 60 | 137 | 150 | [7 empty] | 1
| 1 | 2 | 139.0106 | 136 | 150 | 140 | 142 | 130 | 151 | 142 | 116 | 146 | 130 | 93 | 135 | 50 | 134 | 140 | [7 empty] | 0
| 1 | 3 | 132.9847 | 142 | 145 | 150 | 136 | 110 | 140 | 121 | 129 | 153 | 137 | 99 | 130 | 50 | 133 | 140 | [7 empty] | 16
| 1 | 4 | 134.9636 | 136 | 140 | 140 | 140 | 120 | 151 | 141 | 131 | 189 | 127 | 107 | 135 | 50 | 134 | 140 | [7 empty] | 138.9064
| 1 | 5 | 151.605 | 135 | 155 | 140 | 140 | 110 | 151 | 133 | 132 | 155 | 140 | 110 | 145 | 50 | 138 | 160 | [7 empty] | 115
| 1 | 6 | 148.6734 | 140 | 140 | 130 | 147 | 130 | 152 | 140 | 130 | 151 | 139 | 99 | 125 | 50 | 141 | 150
| Mean |  | 140.6568 | 137.1667 | 145.8333 | 143.3333 | 139.3333 | 120 | 150.3333 | 134.5 | 128.3333 | 154.8333 | 135.6667 | 101.8333 | 134.1667 | 51.6667 | 136.1667 | 146.6667
| Median |  | 137.8571 | 136 | 145 | 140 | 140 | 120 | 151 | 136.5 | 130.5 | 152 | 138 | 101 | 135 | 50 | 135.5 | 145
| Std Dev. |  | 7.6647 | 3.1252 | 5.8452 | 10.328 | 5.4283 | 8.9443 | 5.5737 | 8.167 | 6.1536 | 18.2035 | 5.785 | 6.1455 | 6.6458 | 4.0825 | 3.0605 | 8.165
| Rel.Std.Dev. |  | 5.449% | 2.278% | 4.008% | 7.206% | 3.896% | 7.454% | 3.708% | 6.072% | 4.795% | 11.76% | 4.264% | 6.035% | 4.953% | 7.902% | 2.248% | 5.567%
| PDM3 |  | 1.26% | -1.252% | 4.987% | 3.187% | 0.3073% | -13.61% | 8.226% | -3.172% | -7.612% | 11.47% | -2.332% | -26.69% | -3.412% | -62.8% | -1.972% | 5.587%
| Z-Score (Absolute) |  | 0.43 | 0.05 | 0.99 | 0.72 | 0.29 | 1.79 | 1.69 | 0.23 | 0.63 | 1.22 | 0.11 | 3.75 | 0.27 | 9.33 | 0.05 | 1.08
| Analytical results for Zr in OREAS 905 (Certified Value 270 ppm) | [24 empty] | Certified
Zr | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  | 1
| No | No | 01: 09; 03: 13; 04: 15; 05: 20; 06: 26; 07: 32 |  | ppm
|  |  | 01: BF*MS; 03: BF*OES; 04: PF*OES/MS; 05: PF*MS; 06: ABL*MS; 07: PF*OES/MS |  | 0
0
| 1 | 1 | 01: 300; 03: 254; 04: 269; 05: 247.041; 06: 287; 07: 275 |  | 1
| 1 | 2 | 01: 290; 03: 257; 04: 261; 05: 250.081; 06: 285; 07: 276 |  | 0
| 1 | 3 | 01: 270; 03: 259; 04: 256; 05: 261.574; 06: 290; 07: 276 |  | 16
| 1 | 4 | 01: 320; 03: 257; 04: 270; 05: 253.61; 06: 269; 07: 272 |  | 270.0568
| 1 | 5 | 01: 290; 03: 253; 04: 258; 05: 247.123; 06: 284; 07: 272 |  | 116
| 1 | 6 | 01: 320; 03: 260; 04: 260; 05: 247.183; 06: 271; 07: 267
| Mean |  | 01: 298.3333; 03: 256.6667; 04: 262.3333; 05: 251.102; 06: 281; 07: 273
| Median |  | 01: 295; 03: 257; 04: 260.5; 05: 248.632; 06: 284.5; 07: 273.5
| Std Dev. |  | 01: 19.4079; 03: 2.7325; 04: 5.8195; 05: 5.739; 06: 8.7864; 07: 3.4641
| Rel.Std.Dev. |  | 01: 6.505%; 03: 1.065%; 04: 2.218%; 05: 2.286%; 06: 3.127%; 07: 1.269%
| PDM3 |  | 01: 10.47%; 03: -4.958%; 04: -2.86%; 05: -7.019%; 06: 4.052%; 07: 1.09%
| Z-Score (Absolute) |  | 01: 1.7; 03: 0.61; 04: 0.3; 05: 1.03; 06: 0.74; 07: 0.3
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 905 (Certified Value 0.516 ppm) | [31 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.54 | 0.56941 | 0.55 | 0.6 | 0.51 | 0.52 | 0.49 | 0.55 | 0.41 | 0.6 | 0.5 | 0.49 | < 1 | 0.33 | < 1 | 0.51 | 0.5 | 0.6 | 0.57 | < 0.3 | 0.3 | 0.4 | 0.55 | 0.57 | 0.49 | 0.7 | 0.53 | 0.39 | 0.493 |  | 1
| 1 | 2 | 0.51 | 0.50775 | 0.53 | 0.5 | 0.49 | 0.52 | 0.46 | 0.55 | 0.36 | 0.5 | 0.6 | 0.5 | < 1 | 0.5 | < 1 | 0.56 | 0.4 | 0.59 | 0.46 | < 0.3 | 0.5 | 0.4 | 0.55 | 0.56 | 0.44 | 0.7 | 0.5 | 0.448 | 0.513 |  | 6
| 1 | 3 | 0.53 | 0.5065 | 0.53 | 0.5 | 0.49 | 0.51 | 0.47 | 0.5 | 0.32 | 0.5 | 0.6 | 0.52 | < 1 | 0.51 | < 1 | 0.52 | 0.3 | 0.6 | 0.51 | < 0.3 | 0.6 | 0.5 | 0.45 | 0.57 | 0.45 | 0.7 | 0.55 | 0.498 | 0.518 |  | 16
| 1 | 4 | 0.53 | 0.56403 | 0.53 | 0.5 | 0.51 | 0.53 | 0.45 | 0.55 | 0.27 | 0.6 | 0.6 | 0.49 | < 1 | 0.5 | < 1 | 0.54 | 0.3 | 0.55 | 0.43 | < 0.3 | 0.3 | 0.4 | 0.6 | 0.52 | 0.48 | 0.7 | 0.54 | 0.441 | 0.512 |  | 0.51648
| 1 | 5 | 0.51 | 0.52584 | 0.54 | 0.5 | 0.49 | 0.51 | 0.45 | 0.55 | 0.36 | 0.6 | 0.5 | 0.54 | < 1 | 0.43 | < 1 | 0.54 | 0.4 | 0.6 | 0.44 | 0.3 | 0.3 | 0.5 | 0.55 | 0.54 | 0.44 | 0.8 | 0.52 | 0.484 | 0.494 |  | 118
| 1 | 6 | 0.53 | 0.53375 | 0.53 | 0.5 | 0.49 | 0.51 | 0.42 | 0.5 | 0.29 | 0.5 | 0.6 | 0.66 | < 1 | 0.49 | < 1 | 0.51 | 0.4 | 0.59 | 0.63 | < 0.3 | 0.4 | 0.5 | 0.55 | 0.51 | 0.44 | 0.7 | 0.54 | 0.45 | 0.537
| Mean |  | 0.525 | 0.53455 | 0.535 | 0.51667 | 0.49667 | 0.51667 | 0.45667 | 0.53333 | 0.335 | 0.55 | 0.56667 | 0.53333 |  | 0.46 |  | 0.53 | 0.38333 | 0.58833 | 0.50667 | 0.3 | 0.4 | 0.45 | 0.54167 | 0.545 | 0.45667 | 0.71667 | 0.53 | 0.45183 | 0.51117
| Median |  | 0.53 | 0.5298 | 0.53 | 0.5 | 0.49 | 0.515 | 0.455 | 0.55 | 0.34 | 0.55 | 0.6 | 0.51 |  | 0.495 |  | 0.53 | 0.4 | 0.595 | 0.485 | 0.3 | 0.35 | 0.45 | 0.55 | 0.55 | 0.445 | 0.7 | 0.535 | 0.449 | 0.5125
| Std Dev. |  | 0.012247 | 0.027077 | 0.0083666 | 0.040825 | 0.010328 | 0.008165 | 0.023381 | 0.02582 | 0.051672 | 0.054772 | 0.05164 | 0.065013 |  | 0.069857 |  | 0.02 | 0.075277 | 0.019408 | 0.079666 |  | 0.12649 | 0.054772 | 0.04916 | 0.025884 | 0.022509 | 0.040825 | 0.017889 | 0.037749 | 0.016388
| Rel.Std.Dev. |  | 2.333% | 5.065% | 1.564% | 7.902% | 2.079% | 1.58% | 5.12% | 4.841% | 15.42% | 9.959% | 9.113% | 12.19% |  | 15.19% |  | 3.774% | 19.64% | 3.299% | 15.72% |  | 31.62% | 12.17% | 9.076% | 4.749% | 4.929% | 5.696% | 3.375% | 8.355% | 3.206%
| PDM3 |  | 1.65% | 3.498% | 3.586% | 0.03626% | -3.836% | 0.03626% | -11.58% | 3.263% | -35.14% | 6.49% | 9.717% | 3.263% |  | -10.94% |  | 2.618% | -25.78% | 13.91% | -1.9% | -41.91% | -22.55% | -12.87% | 4.877% | 5.522% | -11.58% | 38.76% | 2.618% | -12.52% | -1.029%
| Z-Score (Absolute) |  | 0.18 | 0.39 | 0.4 | 0 | 0.44 | 0 | 1.32 | 0.37 | 3.99 | 0.73 | 1.1 | 0.19 | NA | 0.67 | NA | 0.29 | 2.93 | 1.74 | 0.22 | 4.76 | 2.57 | 1.47 | 0.95 | 0.62 | 1.32 | 4.4 | 0.29 | 1.43 | 0.12
| Analytical results for Al in OREAS 905 (Certified Value 0.817 wt.%) | [31 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.5g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | 0.84 | 1.0424 | 0.83 | 0.82 | 0.78 | 0.84 | 1.06 | 0.93 | 0.84 | 0.82 | 0.83 | 0.773 | 0.68 | 0.88 | 0.85 | 0.79 | 0.916 | 0.74 | 0.93 | 0.77 | 0.8 | 0.81 | 0.71 | 0.833 | 0.81 | 0.856 | 0.75 |  |  |  | 1
| 1 | 2 | 0.82 | 1.0825 | 0.8 | 0.866 | 0.78 | 0.85 | 1.09 | 0.96 | 0.87 | 0.82 | 0.81 | 0.7954 | 0.65 | 0.84 | 0.83 | 0.77 | 0.9 | 0.78 | 0.87 | 0.79 | 0.79 | 0.8 | 0.72 | 0.833 | 0.79 | 0.817 | 0.75 |  |  |  | 0
| 1 | 3 | 0.85 | 1.0379 | 0.81 | 0.824 | 0.81 | 0.82 | 1.09 | 0.96 | 0.84 | 0.81 | 0.81 | 0.8078 | 0.68 | 0.89 | 0.84 | 0.82 | 0.892 | 0.8 | 0.89 | 0.78 | 0.81 | 0.79 | 0.72 | 0.858 | 0.8 | 0.823 | 0.75 |  |  |  | 16
| 1 | 4 | 0.84 | 1.0012 | 0.79 | 0.835 | 0.78 | 0.84 | 1.09 | 0.97 | 0.87 | 0.84 | 0.83 | 0.8017 | 0.68 | 0.9 | 0.83 | 0.78 | 0.936 | 0.8 | 0.91 | 0.77 | 0.78 | 0.79 | 0.7 | 0.898 | 0.83 | 0.802 | 0.76 |  |  |  | 0.81726
| 1 | 5 | 0.86 | 1.0657 | 0.78 | 0.831 | 0.79 | 0.81 | 1.1 | 0.93 | 0.89 | 0.82 | 0.82 | 0.8409 | 0.68 | 0.9 | 0.85 | 0.78 | 0.916 | 0.81 | 0.92 | 0.78 | 0.76 | 0.78 | 0.69 | 0.84 | 0.82 | 0.816 | 0.75 |  |  |  | 119
| 1 | 6 | 0.85 | 1.0817 | 0.78 | 0.839 | 0.77 | 0.82 | 1.06 | 0.94 | 0.85 | 0.83 | 0.82 | 0.7021 | 0.68 | 0.86 | 0.89 | 0.8 | 0.936 | 0.78 | 0.92 | 0.78 | 0.78 | 0.79 | 0.7 | 0.843 | 0.81 | 0.78 | 0.78
| Mean |  | 0.84333 | 1.0519 | 0.79833 | 0.83583 | 0.785 | 0.83 | 1.0817 | 0.94833 | 0.86 | 0.82333 | 0.82 | 0.78682 | 0.675 | 0.87833 | 0.84833 | 0.79 | 0.916 | 0.785 | 0.90667 | 0.77833 | 0.78667 | 0.79333 | 0.70667 | 0.85083 | 0.81 | 0.81567 | 0.75667
| Median |  | 0.845 | 1.054 | 0.795 | 0.833 | 0.78 | 0.83 | 1.09 | 0.95 | 0.86 | 0.82 | 0.82 | 0.79855 | 0.68 | 0.885 | 0.845 | 0.785 | 0.916 | 0.79 | 0.915 | 0.78 | 0.785 | 0.79 | 0.705 | 0.8415 | 0.81 | 0.8165 | 0.75
| Std Dev. |  | 0.013663 | 0.031214 | 0.019408 | 0.016339 | 0.013784 | 0.015492 | 0.017224 | 0.017224 | 0.02 | 0.010328 | 0.0089443 | 0.046964 | 0.012247 | 0.024014 | 0.022286 | 0.017889 | 0.018067 | 0.0251 | 0.022509 | 0.0075277 | 0.017512 | 0.010328 | 0.012111 | 0.024863 | 0.014142 | 0.025065 | 0.012111
| Rel.Std.Dev. |  | 1.62% | 2.967% | 2.431% | 1.955% | 1.756% | 1.866% | 1.592% | 1.816% | 2.326% | 1.254% | 1.091% | 5.969% | 1.814% | 2.734% | 2.627% | 2.264% | 1.972% | 3.197% | 2.483% | 0.9672% | 2.226% | 1.302% | 1.714% | 2.922% | 1.746% | 3.073% | 1.601%
| PDM3 |  | 3.19% | 28.71% | -2.316% | 2.273% | -3.947% | 1.559% | 32.35% | 16.04% | 5.23% | 0.7432% | 0.3354% | -3.725% | -17.41% | 7.473% | 3.802% | -3.335% | 12.08% | -3.947% | 10.94% | -4.763% | -3.743% | -2.928% | -13.53% | 4.108% | -0.8882% | -0.1949% | -7.414%
| Z-Score (Absolute) |  | 0.47 | 4.69 | 0.44 | 0.2 | 0.81 | 0.2 | 5.29 | 2.6 | 0.81 | 0.07 | 0 | 0.33 | 2.83 | 1.18 | 0.57 | 0.61 | 1.94 | 0.71 | 1.75 | 0.84 | 0.67 | 0.54 | 2.29 | 0.43 | 0.2 | 0.09 | 1.38
| Analytical results for As in OREAS 905 (Certified Value 31.7 ppm) | [31 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 34.7 | 31.411 | 32 | 32 | 31 | 32.7 | 29 | 31.5 | 30 | 37 | 33.8 | 31 | 38 | 30 | 28 | 32.4 | 25 | 32.9 | 32.9 | 31 | 26.7 | 33 | 35.8 | 34.6 | 31.5 | 28 | 31.5 | 33 | 30.9 |  | 1
| 1 | 2 | 32.5 | 30.977 | 31 | 35 | 30.3 | 32.3 | 29 | 32.5 | 31 | 38 | 30.7 | 31 | 32 | 31 | 30 | 32.3 | 25 | 32.7 | 34.3 | 31 | 33.7 | 32 | 34.2 | 33.1 | 28.9 | 29 | 30.9 | 32 | 29.5 |  | 0
| 1 | 3 | 35.1 | 30.794 | 30 | 34 | 31.3 | 32.1 | 28 | 32 | 30 | 38 | 31.1 | 31 | 35 | 31 | 28 | 32.2 | 22 | 32.1 | 34.9 | 32 | 33.9 | 33 | 29 | 33.3 | 29.4 | 29 | 32.9 | 32 | 32.1 |  | 16
| 1 | 4 | 34.2 | 30.5176 | 31 | 35 | 32.2 | 32.6 | 29 | 33 | 31 | 38 | 30.4 | 31 | 41 | 31 | 27 | 31.5 | 20 | 33.3 | 33.4 | 32 | 34.6 | 33 | 38 | 31.6 | 33.1 | 28 | 32.4 | 31 | 32.9 |  | 31.6861
| 1 | 5 | 33.7 | 31.1095 | 31 | 34 | 30.8 | 31.6 | 28 | 33 | 32 | 37 | 31.3 | 32 | 42 | 32 | 30 | 31.4 | 22 | 33.1 | 34.6 | 32 | 34.8 | 32 | 36.8 | 32.1 | 27.5 | 28 | 31.8 | 31 | 32.9 |  | 120
| 1 | 6 | 33.9 | 31.0039 | 31 | 36 | 31.1 | 32.3 | 29 | 32.5 | 31 | 38 | 31.6 | 46 | 38 | 31 | 30 | 31.3 | 27 | 32.8 | 37 | 31 | 34.4 | 33 | 35.8 | 32.1 | 30.1 | 29 | 32.5 | 30 | 33.5
| Mean |  | 34.0167 | 30.9689 | 31 | 34.3333 | 31.1167 | 32.2667 | 28.6667 | 32.4167 | 30.8333 | 37.6667 | 31.4833 | 33.6667 | 37.6667 | 31 | 28.8333 | 31.85 | 23.5 | 32.8167 | 34.5167 | 31.5 | 33.0167 | 32.6667 | 34.9333 | 32.8 | 30.0833 | 28.5 | 32 | 31.5 | 31.9667
| Median |  | 34.05 | 30.9905 | 31 | 34.5 | 31.05 | 32.3 | 29 | 32.5 | 31 | 38 | 31.2 | 31 | 38 | 31 | 29 | 31.85 | 23.5 | 32.85 | 34.45 | 31.5 | 34.15 | 33 | 35.8 | 32.6 | 29.75 | 28.5 | 32.1 | 31.5 | 32.5
| Std Dev. |  | 0.90425 | 0.30028 | 0.63246 | 1.3663 | 0.63061 | 0.39328 | 0.5164 | 0.58452 | 0.75277 | 0.5164 | 1.2123 | 6.0553 | 3.7238 | 0.63246 | 1.3292 | 0.501 | 2.5884 | 0.41191 | 1.4303 | 0.54772 | 3.1224 | 0.5164 | 3.1665 | 1.0954 | 1.9823 | 0.54772 | 0.73756 | 1.0488 | 1.5055
| Rel.Std.Dev. |  | 2.658% | 0.9696% | 2.04% | 3.979% | 2.027% | 1.219% | 1.801% | 1.803% | 2.441% | 1.371% | 3.851% | 17.99% | 9.886% | 2.04% | 4.61% | 1.573% | 11.01% | 1.255% | 4.144% | 1.739% | 9.457% | 1.581% | 9.064% | 3.34% | 6.589% | 1.922% | 2.305% | 3.33% | 4.71%
| PDM3 |  | 7.355% | -2.264% | -2.165% | 8.355% | -1.797% | 1.832% | -9.529% | 2.306% | -2.691% | 18.87% | -0.6399% | 6.251% | 18.87% | -2.165% | -9.003% | 0.5173% | -25.83% | 3.568% | 8.933% | -0.5873% | 4.199% | 3.095% | 10.25% | 3.515% | -5.058% | -10.06% | 0.9907% | -0.5873% | 0.8855%
| Z-Score (Absolute) |  | 1.54 | 0.63 | 0.6 | 1.76 | 0.52 | 0.3 | 2.26 | 0.4 | 0.72 | 4.13 | 0.59 | 0.46 | 4.13 | 0.6 | 2.14 | 0 | 5.93 | 0.69 | 1.89 | 0.25 | 1.72 | 0.58 | 3.03 | 0.67 | 1.25 | 2.38 | 0.11 | 0.25 | 0.08
| Analytical results for Au in OREAS 905 (Certified Value 0.395 ppm) | [31 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 00 | 01 | 03 | 05 | 06 | 07 | 08 | 09 | 12 | 13 | 14 | 17 | 19 | 22 | 26 | 28 | 29 | 30 | 31 | 33 | [10 empty] | ppm
|  |  | INAA | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*MS | AR*MS | AR*AAS | AR*MS | AR*AAS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*MS | [10 empty] | 3
|  |  | 0.09g | 25g | 50g | 25g | 25g | 25g | 50g | 40g | 30g | 25g | 50g | 25g | 30g | 25g | 40g | 25g | 25g | 25g | 25g | 15g | [10 empty] | 3
| 1 | 1 | 0.3796 | 0.41 | 0.37953 | 0.383 | 0.37 | 0.412 | 0.4 | 0.377 | 0.377 | 0.397 | 0.411 | 0.391 | 0.396 | 0.398 | IS | 0.371 | 0.386 | 0.172 | 0.371 | 0.3884 | [10 empty] | 1
| 1 | 2 | 0.4097 | 0.417 | 0.38901 | 0.382 | 0.379 | 0.408 | 0.4 | 0.376 | 0.353 | 0.418 | 0.46 | 0.384 | 0.396 | 0.382 | 0.404 | 0.387 | 0.41 | 0.208 | 0.374 | 0.4111 | [10 empty] | 0
| 1 | 3 | 0.4203 | 0.407 | 0.40142 | 0.391 | 0.371 | 0.409 | 0.4 | 0.378 | 0.345 | 0.426 | 0.454 | 0.379 | 0.381 | 0.388 | 0.399 | 0.37 | 0.427 | 0.2 | 0.381 | 0.3952 | [10 empty] | 16
| 1 | 4 | 0.4304 | 0.406 | 0.40599 | 0.399 | 0.372 | 0.411 | 0.39 | 0.384 | 0.336 | 0.42 | 0.433 | 0.381 | 0.394 | 0.395 | 0.404 | 0.386 | 0.4 | 0.22 | 0.394 | 0.4122 | [10 empty] | 0.39498
| 1 | 5 | 0.3998 | 0.408 | 0.37939 | 0.407 | 0.379 | 0.419 | 0.4 | 0.389 | 0.354 | 0.416 | 0.425 | 0.386 | 0.385 | 0.394 | IS | 0.399 | 0.401 | 0.181 | 0.379 | 0.4308 | [10 empty] | 121
| 1 | 6 | 0.4171 | 0.4 | 0.40187 | 0.393 | 0.385 | 0.413 | 0.39 | 0.373 | 0.396 | 0.488 | 0.423 | 0.378 | 0.391 | 0.398 | 0.395 | 0.402 | 0.392 | 0.157 | 0.385 | 0.4317
|  | 7 | 01: 0.4235
|  | 8 | 01: 0.3908
|  | 9 | 01: 0.4178
|  | 10 | 01: 0.4226
|  | 11 | 01: 0.3792
|  | 12 | 01: 0.3856
|  | 13 | 01: 0.3941
|  | 14 | 01: 0.4024
|  | 15 | 01: 0.4128
|  | 16 | 01: 0.371
|  | 17 | 01: 0.3795
|  | 18 | 01: 0.3988
|  | 19 | 01: 0.4162
|  | 20 | 01: 0.3901
| Mean |  | 0.40207 | 0.408 | 0.39287 | 0.3925 | 0.376 | 0.412 | 0.39667 | 0.3795 | 0.36017 | 0.4275 | 0.43433 | 0.38317 | 0.3905 | 0.3925 | 0.4005 | 0.38583 | 0.40267 | 0.18967 | 0.38067 | 0.41157
| Median |  | 0.4011 | 0.4075 | 0.39521 | 0.392 | 0.3755 | 0.4115 | 0.4 | 0.3775 | 0.3535 | 0.419 | 0.429 | 0.3825 | 0.3925 | 0.3945 | 0.4015 | 0.3865 | 0.4005 | 0.1905 | 0.38 | 0.41165
| Std Dev. |  | 0.017749 | 0.0055498 | 0.011838 | 0.0095446 | 0.005933 | 0.0038987 | 0.005164 | 0.0058907 | 0.02223 | 0.031214 | 0.019012 | 0.0048751 | 0.0062209 | 0.0063166 | 0.0043589 | 0.013467 | 0.014473 | 0.023737 | 0.0082138 | 0.017779
| Rel.Std.Dev. |  | 4.415% | 1.36% | 3.013% | 2.432% | 1.578% | 0.9463% | 1.302% | 1.552% | 6.172% | 7.301% | 4.377% | 1.272% | 1.593% | 1.609% | 1.088% | 3.49% | 3.594% | 12.52% | 2.158% | 4.32%
| PDM3 |  | 1.793% | 3.296% | -0.536% | -0.6287% | -4.806% | 4.308% | 0.4262% | -3.92% | -8.815% | 8.232% | 9.962% | -2.992% | -1.135% | -0.6287% | 1.397% | -2.317% | 1.945% | -51.98% | -3.625% | 4.198%
| Z-Score (Absolute) |  | NA | 0.88 | 0.02 | 0 | 0.94 | 1.11 | 0.24 | 0.74 | 1.84 | 1.3 | 2.09 | 0.53 | 0.11 | 0 | 0.46 | 0.38 | 0.58 | 11.56 | 0.67 | 1.09
Note: Aqua Regia - Au on 10-50g aliquots only.
| Analytical results for B in OREAS 905 (Indicative Value < 10 ppm) | [31 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 12 | 14 | 15 | 17 | 19 | 22 | 23 | 25 | 28 | 29 | 30 | 31 | 32 | 33 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | [10 empty] | 1
|  |  | 0.4g | 20g | 0.25g | 0.5g | 0.5g | 0.25g | 0.5g | 10g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.5g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g | [10 empty] | 1
| 1 | 1 | < 10 | < 10 | < 10 | < 10 | < 10 | 7 | < 10 | < 10 | 60 | < 20 | 35 | < 3 | < 20 | < 20 | < 10 | < 10 | 14 | < 10 | 20.3 | < 20 | [10 empty] | 1
| 1 | 2 | < 10 | < 10 | < 10 | < 10 | < 10 | 12 | < 10 | < 10 | 60 | < 20 | 35 | < 3 | < 20 | < 20 | < 10 | < 10 | 12 | < 10 | 17.2 | < 20 | [10 empty] | 3
| 1 | 3 | < 10 | < 10 | < 10 | < 10 | < 10 | 9 | < 10 | < 10 | 60 | < 20 | 35 | < 3 | < 20 | < 20 | < 10 | < 10 | 11 | < 10 | 16.6 | < 20 | [10 empty] | 16
| 1 | 4 | < 10 | < 10 | < 10 | < 10 | < 10 | 9 | < 10 | < 10 | 60 | < 20 | 35 | < 3 | < 20 | < 20 | < 10 | < 10 | 10 | < 10 | 18.5 | 26 | [10 empty] | < 10
| 1 | 5 | < 10 | < 10 | < 10 | < 10 | < 10 | 17 | < 10 | < 10 | 60 | < 20 | 35 | < 3 | < 20 | < 20 | < 10 | < 10 | 10 | < 10 | 17.4 | < 20 | [10 empty] | 18
| 1 | 6 | < 10 | < 10 | < 10 | < 10 | < 10 | 16 | < 10 | < 10 | 50 | < 20 | 35 | < 3 | < 20 | < 20 | < 10 | < 10 | 8 | < 10 | 17.5 | 23
| Mean |  | 07: 11.6667; 10: 58.3333; 12: 35; 18: 10.8333; 22: 17.9167; 23: 24.5
| Median |  | 07: 10.5; 10: 60; 12: 35; 18: 10.5; 22: 17.45; 23: 24.5
| Std Dev. |  | 07: 4.0825; 10: 4.0825; 12: 0; 18: 2.0412; 22: 1.3197; 23: 2.1213
| Rel.Std.Dev. |  | 07: 34.99%; 10: 6.999%; 12: 0%; 18: 18.84%; 22: 7.366%; 23: 8.658%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | 0.6 | NA | NA | 7.75 | NA | 3.57 | NA | NA | NA | NA | NA | 0.75 | NA | 0.43 | 0.43
| Analytical results for Ba in OREAS 905 (Certified Value 249 ppm) | [31 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 0
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 0
| 1 | 1 | 250 | 304.9343 | 258 | 276 | 240 | 260 | 254 | 244 | 261 | 266 | 254 | 230 | 226 | 264 | 262 | 238.6 | 246 | 160.5 | 223 | 299 | 239 | 255 | 240 | 230 | 256 | 240 | 267 | 224.9 |  |  | 1
| 1 | 2 | 240 | 288.9917 | 249 | 279 | 240 | 260 | 252 | 249 | 275 | 267 | 247 | 231 | 221 | 267 | 242 | 237.4 | 244 | 187 | 233 | 274 | 244 | 271 | 240 | 230 | 256 | 240 | 257 | 223.7 |  |  | 0
| 1 | 3 | 250 | 284.6414 | 250 | 268 | 250 | 250 | 257 | 247 | 260 | 266 | 245 | 238 | 229 | 262 | 234 | 242.8 | 241 | 187 | 239 | 277 | 246 | 247 | 240 | 230 | 256 | 240 | 258 | 219.4 |  |  | 16
| 1 | 4 | 250 | 289.9581 | 253 | 273 | 240 | 260 | 258 | 251 | 278 | 274 | 246 | 235 | 240 | 267 | 224 | 240.3 | 252 | 183 | 239 | 280 | 238 | 276 | 240 | 230 | 262 | 250 | 258 | 227.9 |  |  | 248.984
| 1 | 5 | 260 | 290.5085 | 251 | 269 | 240 | 250 | 260 | 248 | 275 | 265 | 248 | 240 | 226 | 275 | 240 | 237 | 249 | 185 | 238 | 277 | 242 | 267 | 240 | 230 | 256 | 240 | 254 | 213.7 |  |  | 122
| 1 | 6 | 250 | 299.8907 | 255 | 276 | 240 | 250 | 254 | 239 | 270 | 272 | 246 | 215 | 226 | 269 | 243 | 239.5 | 251 | 202.5 | 229 | 282 | 239 | 267 | 240 | 230 | 254 | 240 | 253 | 218.3
| Mean |  | 250 | 293.1541 | 252.6667 | 273.5 | 241.6667 | 255 | 255.8333 | 246.3333 | 269.8333 | 268.3333 | 247.6667 | 231.5 | 228 | 267.3333 | 240.8333 | 239.2667 | 247.1667 | 184.1667 | 233.5 | 281.5 | 241.3333 | 263.8333 | 240 | 230 | 256.6667 | 241.6667 | 257.8333 | 221.3167
| Median |  | 250 | 290.2333 | 252 | 274.5 | 240 | 255 | 255.5 | 247.5 | 272.5 | 266.5 | 246.5 | 233 | 226 | 267 | 241 | 239.05 | 247.5 | 186 | 235.5 | 278.5 | 240.5 | 267 | 240 | 230 | 256 | 240 | 257.5 | 221.55
| Std Dev. |  | 6.3246 | 7.631 | 3.3862 | 4.3243 | 4.0825 | 5.4772 | 2.9944 | 4.274 | 7.679 | 3.7238 | 3.266 | 8.961 | 6.4187 | 4.5019 | 12.5286 | 2.1295 | 4.2622 | 13.5228 | 6.5038 | 9.0056 | 3.2042 | 10.7781 | 0 | 0 | 2.7325 | 4.0825 | 4.9565 | 5.1507
| Rel.Std.Dev. |  | 2.53% | 2.603% | 1.34% | 1.581% | 1.689% | 2.148% | 1.17% | 1.735% | 2.846% | 1.388% | 1.319% | 3.871% | 2.815% | 1.684% | 5.202% | 0.89% | 1.724% | 7.343% | 2.785% | 3.199% | 1.328% | 4.085% | 0% | 0% | 1.065% | 1.689% | 1.922% | 2.327%
| PDM3 |  | 0.4081% | 17.74% | 1.479% | 9.846% | -2.939% | 2.416% | 2.751% | -1.065% | 8.374% | 7.771% | -0.5291% | -7.022% | -8.428% | 7.37% | -3.274% | -3.903% | -0.7299% | -26.03% | -6.219% | 13.06% | -3.073% | 5.964% | -3.608% | -7.625% | 3.086% | -2.939% | 3.554% | -11.11%
| Z-Score (Absolute) |  | 0.23 | 3.27 | 0.42 | 1.89 | 0.36 | 0.58 | 0.64 | 0.03 | 1.63 | 1.52 | 0.03 | 1.08 | 1.5 | 1.45 | 0.42 | 0.53 | 0.03 | 4.09 | 0.94 | 2.2 | 0.38 | 1.2 | 0.48 | 1.19 | 0.7 | 0.36 | 0.65 | 1.8
| Analytical results for Be in OREAS 905 (Certified Value 0.93 ppm) | [31 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 24 | 28 | 29 | 30 | 31 | 32 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.4g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [6 empty] | 3
| 1 | 1 | 1.04 | 0.97057 | 1 | 0.9 | 0.83 | 0.93 | 0.92 | 0.9 | 1 | 1 | 0.9 | 1 | 0.9 | 0.9 | 0.98 | 0.8 | 0.9 | 0.72 | < 1 | 1 | 1 | 0.9 | 0.85 | 1 | [6 empty] | 1
| 1 | 2 | 0.97 | 1.0504 | 1.1 | 0.9 | 0.81 | 0.92 | 0.96 | 1 | 1 | 0.9 | 0.8 | < 1 | 0.9 | 0.9 | 0.96 | 0.8 | 0.9 | 0.78 | 1 | 1.01 | 0.92 | 0.9 | 0.88 | 0.9 | [6 empty] | 0
| 1 | 3 | 1.01 | 1.1109 | 1 | 0.9 | 0.83 | 0.89 | 0.95 | 0.9 | 1 | 0.9 | 0.9 | 1 | 0.9 | 0.9 | 0.92 | 0.8 | 0.9 | 0.8 | 1 | 1.04 | 0.99 | 0.8 | 0.87 | 1 | [6 empty] | 16
| 1 | 4 | 1.01 | 1.1078 | 1 | 0.9 | 0.88 | 0.91 | 0.98 | 1 | 1 | 1 | 0.9 | 1 | 1 | 0.8 | 0.96 | 0.8 | 0.9 | 0.76 | 1 | 0.88 | 1.02 | 0.9 | 0.92 | 1 | [6 empty] | 0.93372
| 1 | 5 | 1.02 | 0.97071 | 0.9 | 1 | 0.82 | 0.84 | 0.94 | 1 | 1 | 1 | 1 | 1 | 0.9 | 0.9 | 0.96 | 0.8 | 0.9 | 0.78 | 1 | 0.95 | 0.95 | 0.8 | 0.86 | 1 | [6 empty] | 123
| 1 | 6 | 1.04 | 0.98002 | 1.1 | 0.9 | 0.85 | 0.89 | 0.96 | 1 | 1 | 1 | 0.9 | 1 | 1 | 0.9 | 1 | 0.8 | 0.9 | 0.81 | 1 | 0.93 | 0.79 | 0.8 | 0.89 | 1
| Mean |  | 1.015 | 1.0317 | 1.0167 | 0.91667 | 0.83667 | 0.89667 | 0.95167 | 0.96667 | 1 | 0.96667 | 0.9 | 1 | 0.93333 | 0.88333 | 0.96333 | 0.8 | 0.9 | 0.775 | 1 | 0.96833 | 0.945 | 0.85 | 0.87833 | 0.98333
| Median |  | 1.015 | 1.0152 | 1 | 0.9 | 0.83 | 0.9 | 0.955 | 1 | 1 | 1 | 0.9 | 1 | 0.9 | 0.9 | 0.96 | 0.8 | 0.9 | 0.78 | 1 | 0.975 | 0.97 | 0.85 | 0.875 | 1
| Std Dev. |  | 0.025884 | 0.067133 | 0.075277 | 0.040825 | 0.025033 | 0.032042 | 0.020412 | 0.05164 | 0 | 0.05164 | 0.063246 | 0 | 0.05164 | 0.040825 | 0.026583 | 1.2162e-16 | 0 | 0.032094 | 0 | 0.059133 | 0.084083 | 0.054772 | 0.024833 | 0.040825
| Rel.Std.Dev. |  | 2.55% | 6.507% | 7.404% | 4.454% | 2.992% | 3.573% | 2.145% | 5.342% | 0% | 5.342% | 7.027% | 0% | 5.533% | 4.622% | 2.76% | 1.52e-14% | 0% | 4.141% | 0% | 6.107% | 8.898% | 6.444% | 2.827% | 4.152%
| PDM3 |  | 8.704% | 10.5% | 8.883% | -1.827% | -10.39% | -3.969% | 1.922% | 3.528% | 7.098% | 3.528% | -3.612% | 7.098% | -0.04192% | -5.397% | 3.171% | -14.32% | -3.612% | -17% | 7.098% | 3.707% | 1.208% | -8.967% | -5.932% | 5.313%
| Z-Score (Absolute) |  | 0.78 | 1 | 0.8 | 0.55 | 1.63 | 0.82 | 0.08 | 0.12 | 0.57 | 0.12 | 0.78 | 0.57 | 0.33 | 1 | 0.08 | 2.12 | 0.78 | 2.46 | 0.57 | 0.15 | 0.25 | 1.45 | 1.07 | 0.35
| Analytical results for Bi in OREAS 905 (Certified Value 5.74 ppm) | [31 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 5.92 | 5.6776 | 6.08 | 5.8 | 5.13 | 5.94 | 5.37 | 5.6 | 5.49 | 5 | 5.8 | 5.73 | < 10 | 5.86 | < 5 | 6.08 | 2 | 8.9 | 4.53 | 6 | 6.4 | 4 | 6.28 | 6.25 | 6.02 | < 5 | 5.22 | 5.71 | 6.47 |  | 1
| 1 | 2 | 5.71 | 5.5929 | 5.83 | 5.8 | 5.04 | 5.92 | 5.54 | 5.5 | 5.58 | 5 | 5.67 | 5.96 | < 10 | 5.89 | < 5 | 5.95 | 2 | 8.3 | 3.68 | 5 | 6.9 | 5 | 6.06 | 6.28 | 5.84 | < 5 | 5.18 | 5.67 | 6.69 |  | 0
| 1 | 3 | 6.01 | 5.5858 | 5.92 | 5.9 | 5.28 | 6.1 | 5.34 | 5.8 | 5.47 | 5 | 5.7 | 6.04 | < 10 | 5.91 | < 5 | 6.18 | 3 | 9.5 | 4.46 | 5 | 7 | 4 | 5.96 | 6.26 | 6.08 | < 5 | 5.47 | 5.56 | 5.94 |  | 16
| 1 | 4 | 5.95 | 5.6416 | 5.83 | 5.8 | 5.34 | 6.29 | 5.41 | 5.5 | 5.53 | 5 | 5.79 | 5.81 | < 10 | 5.92 | < 5 | 5.96 | 3 | 7.8 | 4.32 | 3 | 6.9 | 4 | 6.08 | 6.2 | 6.16 | < 5 | 5.5 | 5.81 | 6.45 |  | 5.7447
| 1 | 5 | 6.01 | 5.65 | 5.94 | 6 | 4.99 | 5.88 | 5.49 | 5.5 | 5.64 | 5 | 5.64 | 5.97 | < 10 | 5.57 | < 5 | 5.96 | 2 | 7.9 | 3.98 | 4 | 7 | 3 | 6.36 | 6.04 | 5.26 | < 5 | 5.51 | 5.6 | 5.76 |  | 124
| 1 | 6 | 5.83 | 5.5862 | 6.02 | 5.7 | 5.1 | 5.97 | 5.49 | 5.5 | 5.56 | 5 | 5.68 | 6.06 | < 10 | 5.77 | < 5 | 6.04 | 5 | 9.6 | 5.69 | 4 | 7 | 4 | 5.94 | 6.12 | 4.54 | < 5 | 5.41 | 5.7 | 5.95
| Mean |  | 5.905 | 5.6224 | 5.9367 | 5.8333 | 5.1467 | 6.0167 | 5.44 | 5.5667 | 5.545 | 5 | 5.7133 | 5.9283 |  | 5.82 |  | 6.0283 | 2.8333 | 8.6667 | 4.4433 | 4.5 | 6.8667 | 4 | 6.1133 | 6.1917 | 5.65 |  | 5.3817 | 5.675 | 6.21
| Median |  | 5.935 | 5.6172 | 5.93 | 5.8 | 5.115 | 5.955 | 5.45 | 5.5 | 5.545 | 5 | 5.69 | 5.965 |  | 5.875 |  | 6 | 2.5 | 8.6 | 4.39 | 4.5 | 6.95 | 4 | 6.07 | 6.225 | 5.93 |  | 5.44 | 5.685 | 6.2
| Std Dev. |  | 0.11658 | 0.039236 | 0.10053 | 0.10328 | 0.13677 | 0.15345 | 0.078486 | 0.12111 | 0.062209 | 0 | 0.066232 | 0.13106 |  | 0.13387 |  | 0.090866 | 1.169 | 0.78655 | 0.68913 | 1.0488 | 0.23381 | 0.63246 | 0.17096 | 0.093897 | 0.6328 |  | 0.14552 | 0.088261 | 0.37379
| Rel.Std.Dev. |  | 1.974% | 0.6979% | 1.693% | 1.771% | 2.657% | 2.55% | 1.443% | 2.176% | 1.122% | 0% | 1.159% | 2.211% |  | 2.3% |  | 1.507% | 41.26% | 9.076% | 15.51% | 23.31% | 3.405% | 15.81% | 2.796% | 1.517% | 11.2% |  | 2.704% | 1.555% | 6.019%
| PDM3 |  | 2.79% | -2.13% | 3.341% | 1.542% | -10.41% | 4.734% | -5.305% | -3.1% | -3.477% | -12.96% | -0.5466% | 3.196% |  | 1.31% |  | 4.937% | -50.68% | 50.86% | -22.65% | -21.67% | 19.53% | -30.37% | 6.416% | 7.78% | -1.649% |  | -6.32% | -1.214% | 8.099%
| Z-Score (Absolute) |  | 0.3 | 0.35 | 0.37 | 0.14 | 1.44 | 0.43 | 0.76 | 0.58 | 0.52 | 1.77 | 0.14 | 0.36 | NA | 0.22 | NA | 0.58 | 6.74 | 6.63 | 3.05 | 2.92 | 2.72 | 4.06 | 0.78 | 0.96 | 0.23 | NA | 0.9 | 0.23 | 1
| Analytical results for Ca in OREAS 905 (Certified Value 0.337 wt.%) | [31 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.34 | 0.32413 | 0.34 | 0.334 | 0.35 | 0.33 | 0.35 | 0.41 | 0.32 | 0.37 | 0.36 | 0.33 | 0.33 | 0.33 | 0.36 | 0.33 | 0.3 | 0.31 | 0.3195 | 0.33 | 0.39 | 0.32 | 0.36 | 0.35 | 0.34 | 0.369 | 0.34 | 0.308 | 0.31 |  | 1
| 1 | 2 | 0.33 | 0.31538 | 0.33 | 0.336 | 0.34 | 0.33 | 0.35 | 0.41 | 0.34 | 0.37 | 0.35 | 0.34 | 0.32 | 0.32 | 0.36 | 0.32 | 0.3 | 0.306 | 0.344 | 0.35 | 0.37 | 0.33 | 0.34 | 0.35 | 0.34 | 0.365 | 0.33 | 0.292 | 0.31 |  | 0
| 1 | 3 | 0.34 | 0.32162 | 0.33 | 0.334 | 0.34 | 0.33 | 0.35 | 0.41 | 0.33 | 0.37 | 0.35 | 0.33 | 0.33 | 0.33 | 0.37 | 0.34 | 0.29 | 0.306 | 0.348 | 0.35 | 0.38 | 0.32 | 0.34 | 0.35 | 0.34 | 0.364 | 0.34 | 0.293 | 0.32 |  | 16
| 1 | 4 | 0.34 | 0.31825 | 0.32 | 0.342 | 0.33 | 0.33 | 0.34 | 0.42 | 0.34 | 0.37 | 0.36 | 0.33 | 0.33 | 0.33 | 0.35 | 0.33 | 0.3 | 0.315 | 0.3305 | 0.35 | 0.39 | 0.32 | 0.33 | 0.35 | 0.34 | 0.363 | 0.34 | 0.293 | 0.31 |  | 0.33678
| 1 | 5 | 0.35 | 0.32338 | 0.32 | 0.349 | 0.34 | 0.32 | 0.35 | 0.41 | 0.33 | 0.37 | 0.36 | 0.34 | 0.33 | 0.33 | 0.38 | 0.33 | 0.3 | 0.315 | 0.3365 | 0.36 | 0.39 | 0.32 | 0.36 | 0.35 | 0.34 | 0.373 | 0.34 | 0.298 | 0.31 |  | 125
| 1 | 6 | 0.34 | 0.32273 | 0.31 | 0.34 | 0.34 | 0.32 | 0.34 | 0.41 | 0.33 | 0.38 | 0.35 | 0.33 | 0.33 | 0.33 | 0.36 | 0.33 | 0.31 | 0.31 | 0.359 | 0.34 | 0.39 | 0.33 | 0.35 | 0.35 | 0.35 | 0.363 | 0.34 | 0.291 | 0.32
| Mean |  | 0.34 | 0.32092 | 0.325 | 0.33917 | 0.34 | 0.32667 | 0.34667 | 0.41167 | 0.33167 | 0.37167 | 0.355 | 0.33333 | 0.32833 | 0.32833 | 0.36333 | 0.33 | 0.3 | 0.31033 | 0.33958 | 0.34667 | 0.385 | 0.32333 | 0.34667 | 0.35 | 0.34167 | 0.36617 | 0.33833 | 0.29583 | 0.31333
| Median |  | 0.34 | 0.32217 | 0.325 | 0.338 | 0.34 | 0.33 | 0.35 | 0.41 | 0.33 | 0.37 | 0.355 | 0.33 | 0.33 | 0.33 | 0.36 | 0.33 | 0.3 | 0.31 | 0.34025 | 0.35 | 0.39 | 0.32 | 0.345 | 0.35 | 0.34 | 0.3645 | 0.34 | 0.293 | 0.31
| Std Dev. |  | 0.0063246 | 0.0034062 | 0.010488 | 0.0058109 | 0.0063246 | 0.005164 | 0.005164 | 0.0040825 | 0.0075277 | 0.0040825 | 0.0054772 | 0.005164 | 0.0040825 | 0.0040825 | 0.010328 | 0.0063246 | 0.0063246 | 0.0040332 | 0.013876 | 0.010328 | 0.0083666 | 0.005164 | 0.012111 | 6.0809e-17 | 0.0040825 | 0.0040208 | 0.0040825 | 0.0064317 | 0.005164
| Rel.Std.Dev. |  | 1.86% | 1.061% | 3.227% | 1.713% | 1.86% | 1.581% | 1.49% | 0.9917% | 2.27% | 1.098% | 1.543% | 1.549% | 1.243% | 1.243% | 2.843% | 1.917% | 2.108% | 1.3% | 4.086% | 2.979% | 2.173% | 1.597% | 3.493% | 1.737e-14% | 1.195% | 1.098% | 1.207% | 2.174% | 1.648%
| PDM3 |  | 0.9574% | -4.71% | -3.497% | 0.71% | 0.9574% | -3.002% | 2.937% | 22.24% | -1.517% | 10.36% | 5.411% | -1.022% | -2.507% | -2.507% | 7.886% | -2.012% | -10.92% | -7.852% | 0.8337% | 2.937% | 14.32% | -3.991% | 2.937% | 3.927% | 1.452% | 8.727% | 0.4625% | -12.16% | -6.961%
| Z-Score (Absolute) |  | 0.05 | 1.14 | 0.88 | 0 | 0.05 | 0.78 | 0.47 | 4.51 | 0.47 | 2.02 | 0.99 | 0.36 | 0.67 | 0.67 | 1.5 | 0.57 | 2.44 | 1.79 | 0.03 | 0.47 | 2.85 | 0.99 | 0.47 | 0.67 | 0.16 | 1.68 | 0.05 | 2.85 | 1.61
| Analytical results for Cd in OREAS 905 (Certified Value 0.34 ppm) | [31 empty] | Certified
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  |  | 3
| 1 | 1 | 0.37 | 0.37351 | 0.39 | < 1 | 0.34 | 0.35 | 0.33 | 0.3 | 0.45 | < 1 | 0.39 | 0.36 | < 1 | 0.33 | 0.35 | 0.69 | 0.32 | < 0.5 | 0.3 | < 0.5 | 0.3 | 0.39 | 0.39 | 0.3 | 0.35 | 0.35 | 0.32 |  |  |  | 1
| 1 | 2 | 0.32 | 0.32923 | 0.36 | < 1 | 0.3 | 0.36 | 0.32 | 0.3 | 0.45 | < 1 | 0.47 | 0.37 | < 1 | 0.33 | 0.35 | 0.74 | 0.35 | < 0.5 | 0.5 | < 0.5 | 0.3 | 0.33 | 0.33 | 0.4 | 0.33 | 0.34 | 0.35 |  |  |  | 8
| 1 | 3 | 0.35 | 0.35393 | 0.35 | < 1 | 0.32 | 0.35 | 0.33 | 0.3 | 0.42 | < 1 | 0.33 | 0.37 | < 1 | 0.33 | 0.35 | 0.67 | 0.37 | < 0.5 | 0.5 | < 0.5 | 0.25 | 0.36 | 0.34 | 0.4 | 0.35 | 0.34 | 0.36 |  |  |  | 16
| 1 | 4 | 0.35 | 0.37487 | 0.37 | < 1 | 0.32 | 0.34 | 0.28 | 0.3 | 0.42 | < 1 | 0.29 | 0.36 | < 1 | 0.34 | 0.36 | 0.67 | 0.35 | < 0.5 | 0.5 | < 0.5 | 0.35 | 0.36 | 0.38 | 0.4 | 0.36 | 0.34 | 0.35 |  |  |  | 0.34212
| 1 | 5 | 0.34 | 0.39668 | 0.35 | < 1 | 0.3 | 0.35 | 0.31 | 0.3 | 0.43 | < 1 | 0.23 | 0.37 | < 1 | 0.31 | 0.36 | 0.71 | 0.34 | < 0.5 | 0.5 | < 0.5 | 0.3 | 0.33 | 0.32 | 0.4 | 0.34 | 0.34 | 0.34 |  |  |  | 126
| 1 | 6 | 0.35 | 0.35306 | 0.35 | < 1 | 0.32 | 0.35 | 0.29 | 0.3 | 0.41 | < 1 | 0.54 | 0.45 | < 1 | 0.34 | 0.31 | 0.68 | 0.38 | < 0.5 | 0.4 | < 0.5 | 0.35 | 0.33 | 0.34 | 0.3 | 0.34 | 0.35 | 0.35
| Mean |  | 0.34667 | 0.36354 | 0.36167 |  | 0.31667 | 0.35 | 0.31 | 0.3 | 0.43 |  | 0.375 | 0.38 |  | 0.33 | 0.34667 | 0.69333 | 0.35167 |  | 0.45 |  | 0.30833 | 0.35 | 0.35 | 0.36667 | 0.345 | 0.34333 | 0.345
| Median |  | 0.35 | 0.36372 | 0.355 |  | 0.32 | 0.35 | 0.315 | 0.3 | 0.425 |  | 0.36 | 0.37 |  | 0.33 | 0.35 | 0.685 | 0.35 |  | 0.5 |  | 0.3 | 0.345 | 0.34 | 0.4 | 0.345 | 0.34 | 0.35
| Std Dev. |  | 0.01633 | 0.023261 | 0.016021 |  | 0.015055 | 0.0063246 | 0.020976 | 0 | 0.016733 |  | 0.11554 | 0.034641 |  | 0.010954 | 0.018619 | 0.027325 | 0.02137 |  | 0.083666 |  | 0.037639 | 0.024495 | 0.028284 | 0.05164 | 0.010488 | 0.005164 | 0.013784
| Rel.Std.Dev. |  | 4.711% | 6.399% | 4.43% |  | 4.754% | 1.807% | 6.767% | 0% | 3.891% |  | 30.81% | 9.116% |  | 3.32% | 5.371% | 3.941% | 6.077% |  | 18.59% |  | 12.21% | 6.999% | 8.081% | 14.08% | 3.04% | 1.504% | 3.995%
| PDM3 |  | 1.33% | 6.263% | 5.715% |  | -7.439% | 2.304% | -9.388% | -12.31% | 25.69% |  | 9.612% | 11.07% |  | -3.542% | 1.33% | 102.7% | 2.792% |  | 31.53% |  | -9.875% | 2.304% | 2.304% | 7.176% | 0.8429% | 0.3557% | 0.8429%
| Z-Score (Absolute) |  | 0.17 | 0.68 | 0.59 | NA | 1.67 | 0 | 2.01 | 2.51 | 4.01 | NA | 0.67 | 0.8 | NA | 1 | 0.2 | 17.23 | 0.08 | NA | 5.02 | NA | 2.09 | 0 | 0 | 0.84 | 0.25 | 0.33 | 0
| Analytical results for Ce in OREAS 905 (Certified Value 80 ppm) | [31 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 0
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 0.25g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [8 empty] | 1
| 1 | 1 | 80.9 | 83.9676 | 78.98 | 77.6 | 74.3 | 83.3 | 81.5 | 77.2 | 69.7 | 85 | 76 | 74.12 | 76.8 | 81.22 | 106.9 | 82 | 81.3 | 87.2 | 89.2 | 58 | 82.3 | 83 | [8 empty] | 1
| 1 | 2 | 73.3 | 81.5565 | 76.52 | 78.8 | 73.3 | 82.5 | 81 | 77.2 | 71.3 | 85 | 73.7 | 76.39 | 77.7 | 80.3 | 116.3 | 92 | 75.2 | 81 | 87.2 | 55.6 | 74.6 | 81.7 | [8 empty] | 0
| 1 | 3 | 79.4 | 82.8449 | 77.54 | 75.9 | 74 | 82.2 | 82.9 | 78 | 68.2 | 85 | 75.4 | 77.31 | 78.1 | 82.62 | 117.2 | 93 | 67.3 | 78.2 | 91.4 | 60.4 | 79.7 | 83.8 | [8 empty] | 16
| 1 | 4 | 78.9 | 83.2437 | 74.77 | 75.6 | 74.8 | 82 | 82.2 | 76 | 68.7 | 86 | 74.6 | 75.41 | 78.1 | 79.11 | 112.9 | 89 | 83.1 | 74.7 | 94.2 | 57.4 | 79.6 | 84.6 | [8 empty] | 79.5606
| 1 | 5 | 77.6 | 82.3508 | 77.5 | 75 | 72.3 | 80.5 | 82.2 | 76.8 | 70.9 | 86 | 76 | 76.21 | 78.5 | 78.8 | 114.4 | 89 | 84.3 | 78.6 | 83.2 | 57.4 | 76.7 | 82.7 | [8 empty] | 127
| 1 | 6 | 79.4 | 83.506 | 78.37 | 77.9 | 72.9 | 82.4 | 81.7 | 76.2 | 67.7 | 86 | 75.3 | 72.58 | 74.5 | 80.33 | 121.1 | 90 | 78.8 | 75.9 | 69.9 | 58.1 | 77.4 | 79.6
| Mean |  | 78.25 | 82.9116 | 77.28 | 76.8 | 73.6 | 82.15 | 81.9167 | 76.9 | 69.4167 | 85.5 | 75.1667 | 75.3367 | 77.2833 | 80.3967 | 114.8 | 89.1667 | 78.3333 | 79.2667 | 85.85 | 57.8167 | 78.3833 | 82.5667
| Median |  | 79.15 | 83.0443 | 77.52 | 76.75 | 73.65 | 82.3 | 81.95 | 77 | 69.2 | 85.5 | 75.35 | 75.81 | 77.9 | 80.315 | 115.35 | 89.5 | 80.05 | 78.4 | 88.2 | 57.7 | 78.5 | 82.85
| Std Dev. |  | 2.6463 | 0.86418 | 1.4881 | 1.506 | 0.93381 | 0.9225 | 0.66156 | 0.73485 | 1.4675 | 0.54772 | 0.88694 | 1.7234 | 1.4811 | 1.4028 | 4.7724 | 3.8687 | 6.307 | 4.4657 | 8.6602 | 1.5523 | 2.7081 | 1.7558
| Rel.Std.Dev. |  | 3.382% | 1.042% | 1.926% | 1.961% | 1.269% | 1.123% | 0.8076% | 0.9556% | 2.114% | 0.6406% | 1.18% | 2.288% | 1.916% | 1.745% | 4.157% | 4.339% | 8.052% | 5.634% | 10.09% | 2.685% | 3.455% | 2.126%
| PDM3 |  | -1.647% | 4.212% | -2.866% | -3.47% | -7.492% | 3.255% | 2.961% | -3.344% | -12.75% | 7.465% | -5.523% | -5.309% | -2.862% | 1.051% | 44.29% | 12.07% | -1.543% | -0.3694% | 7.905% | -27.33% | -1.48% | 3.778%
| Z-Score (Absolute) |  | 0.08 | 0.81 | 0.3 | 0.4 | 1.03 | 0.66 | 0.61 | 0.38 | 1.86 | 1.32 | 0.72 | 0.69 | 0.19 | 0.31 | 7.12 | 2.05 | 0.09 | 0.09 | 2.02 | 4.16 | 0.08 | 0.74
| Analytical results for Co in OREAS 905 (Certified Value 13.9 ppm) | [31 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 15.5 | 13.8456 | 14.5 | 14 | 13.5 | 14.9 | 13 | 14.4 | 13.5 | 15 | 14 | 13.5 | 13 | 13.8 | 14 | 13.64 | 12.3 | 14.5 | 11.4 | 14 | 17.8 | 13 | 14 | 14.8 | 14 | 14 | 13 | 14.7 | 14.7 |  | 1
| 1 | 2 | 13.5 | 13.5896 | 13.5 | 14 | 13.4 | 14.6 | 14 | 14.6 | 13.7 | 15 | 14 | 13.3 | 12 | 13.7 | 14 | 13.39 | 11.9 | 14.5 | 12.3 | 14 | 15.6 | 13 | 14 | 14.3 | 14.9 | 14 | 13 | 14.1 | 14 |  | 0
| 1 | 3 | 15.5 | 14.2426 | 13.7 | 14 | 13.6 | 14.2 | 13 | 14.4 | 12.9 | 15 | 14 | 13.6 | 13 | 14.2 | 13 | 13.69 | 11.2 | 14.2 | 12.5 | 14 | 15.8 | 13 | 15 | 14.3 | 15.6 | 13 | 13.5 | 14.3 | 14.5 |  | 16
| 1 | 4 | 15 | 13.8445 | 13.9 | 14 | 14.1 | 14.4 | 13 | 14.6 | 13.2 | 15 | 14 | 13.8 | 13 | 13.8 | 13 | 13.45 | 11.6 | 14.7 | 12 | 15 | 16.1 | 13 | 15 | 13.6 | 15.7 | 14 | 13.5 | 14.6 | 14.5 |  | 13.8923
| 1 | 5 | 14.9 | 13.9441 | 13.7 | 14 | 13.5 | 14 | 14 | 14.2 | 13.5 | 15 | 14 | 13.9 | 13 | 13.8 | 13 | 13.27 | 11.7 | 14.4 | 12.3 | 15 | 15.9 | 13 | 16 | 13.9 | 14.2 | 14 | 13.4 | 14.6 | 14.9 |  | 128
| 1 | 6 | 14.1 | 13.5982 | 13.8 | 14 | 13.6 | 14.3 | 14 | 14.6 | 12.8 | 15 | 14 | 14.1 | 14 | 13.7 | 13 | 13.24 | 11.6 | 14.5 | 12.6 | 14 | 16 | 13 | 14 | 14 | 13.1 | 14 | 13.5 | 14 | 14.9
| Mean |  | 14.75 | 13.8441 | 13.85 | 14 | 13.6167 | 14.4 | 13.5 | 14.4667 | 13.2667 | 15 | 14 | 13.7 | 13 | 13.8333 | 13.3333 | 13.4467 | 11.7167 | 14.4667 | 12.1833 | 14.3333 | 16.2 | 13 | 14.6667 | 14.15 | 14.5833 | 13.8333 | 13.3167 | 14.3833 | 14.5833
| Median |  | 14.95 | 13.845 | 13.75 | 14 | 13.55 | 14.35 | 13.5 | 14.5 | 13.35 | 15 | 14 | 13.7 | 13 | 13.8 | 13 | 13.42 | 11.65 | 14.5 | 12.3 | 14 | 15.95 | 13 | 14.5 | 14.15 | 14.55 | 14 | 13.45 | 14.45 | 14.6
| Std Dev. |  | 0.79937 | 0.24249 | 0.34496 | 0 | 0.24833 | 0.31623 | 0.54772 | 0.1633 | 0.36148 | 0 | 0 | 0.28983 | 0.63246 | 0.18619 | 0.5164 | 0.1864 | 0.3656 | 0.1633 | 0.43551 | 0.5164 | 0.8025 | 0 | 0.8165 | 0.41352 | 1.0068 | 0.40825 | 0.24833 | 0.29269 | 0.33714
| Rel.Std.Dev. |  | 5.419% | 1.752% | 2.491% | 0% | 1.824% | 2.196% | 4.057% | 1.129% | 2.725% | 0% | 0% | 2.116% | 4.865% | 1.346% | 3.873% | 1.386% | 3.12% | 1.129% | 3.575% | 3.603% | 4.954% | 0% | 5.567% | 2.922% | 6.904% | 2.951% | 1.865% | 2.035% | 2.312%
| PDM3 |  | 6.174% | -0.3472% | -0.3047% | 0.7751% | -1.984% | 3.654% | -2.824% | 4.134% | -4.504% | 7.973% | 0.7751% | -1.384% | -6.423% | -0.4246% | -4.024% | -3.208% | -15.66% | 4.134% | -12.3% | 3.174% | 16.61% | -6.423% | 5.574% | 1.855% | 4.974% | -0.4246% | -4.144% | 3.534% | 4.974%
| Z-Score (Absolute) |  | 1.13 | 0 | 0.16 | 0.19 | 0.41 | 0.7 | 0.43 | 0.78 | 0.72 | 1.45 | 0.19 | 0.18 | 1.06 | 0.01 | 0.64 | 0.5 | 2.66 | 0.78 | 2.08 | 0.61 | 2.55 | 1.06 | 1.03 | 0.38 | 0.92 | 0.01 | 0.66 | 0.67 | 0.92
| Analytical results for Cr in OREAS 905 (Certified Value 17.6 ppm) | [31 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 1
| 1 | 1 | 18 | 18.5956 | 21 | 21 | 17 | 17 | 18 | 20 | 17 | < 10000 | 19 | 17 | 17 | 22 | 49 | 19 | 13.8 | 22.9 | 26.085 | 17 | 20 | 16 | 15 | 18 | 17 | 16 | 18 | 17.4 | 17 |  | 1
| 1 | 2 | 17 | 18.8039 | 19 | 21 | 17 | 18 | 17 | 20 | 18 | < 10000 | 16 | 17 | 14 | 19 | 49 | 17 | 13.9 | 23.4 | 21.83 | 18 | 22 | 16 | 15 | 19 | 17 | 17 | 16 | 17.5 | 17.7 |  | 9
| 1 | 3 | 17 | 18.0856 | 19 | 20 | 17 | 18 | 17 | 20 | 17 | < 10000 | 17 | 17 | 15 | 19 | 29 | 20 | 13.2 | 23.4 | 22.2925 | 17 | 24 | 17 | 15 | 17 | 17 | 15 | 17 | 16.9 | 18.4 |  | 16
| 1 | 4 | 18 | 17.9841 | 20 | 21 | 18 | 18 | 17 | 20 | 19 | < 10000 | 19 | 17 | 14 | 20 | 43 | 19 | 13.7 | 23 | 20.9975 | 18 | 21 | 16 | 15 | 17 | 18 | 15 | 17 | 17.5 | 17.2 |  | 17.6185
| 1 | 5 | 18 | 18.6936 | 20 | 21 | 17 | 17 | 17 | 20 | 18 | < 10000 | 17 | 17 | 14 | 18 | 34 | 17 | 15.6 | 22.4 | 21.275 | 17 | 21 | 17 | 15 | 17 | 17 | 16 | 18 | 19.5 | 18.1 |  | 129
| 1 | 6 | 18 | 18.5386 | 20 | 21 | 18 | 17 | 17 | 20 | 18 | < 10000 | 18 | 16 | 15 | 18 | 44 | 18 | 15.1 | 22.8 | 23.2175 | 20 | 21 | 16 | 20 | 18 | 16 | 15 | 16 | 16.6 | 19.3
| Mean |  | 17.6667 | 18.4502 | 19.8333 | 20.8333 | 17.3333 | 17.5 | 17.1667 | 20 | 17.8333 |  | 17.6667 | 16.8333 | 14.8333 | 19.3333 | 41.3333 | 18.3333 | 14.2167 | 22.9833 | 22.6163 | 17.8333 | 21.5 | 16.3333 | 15.8333 | 17.6667 | 17 | 15.6667 | 17 | 17.5667 | 17.95
| Median |  | 18 | 18.5671 | 20 | 21 | 17 | 17.5 | 17 | 20 | 18 |  | 17.5 | 17 | 14.5 | 19 | 43.5 | 18.5 | 13.85 | 22.95 | 22.0613 | 17.5 | 21 | 16 | 15 | 17.5 | 17 | 15.5 | 17 | 17.45 | 17.9
| Std Dev. |  | 0.5164 | 0.33571 | 0.75277 | 0.40825 | 0.5164 | 0.54772 | 0.40825 | 0 | 0.75277 |  | 1.2111 | 0.40825 | 1.169 | 1.5055 | 8.165 | 1.2111 | 0.92394 | 0.38166 | 1.8728 | 1.169 | 1.3784 | 0.5164 | 2.0412 | 0.8165 | 0.63246 | 0.8165 | 0.89443 | 1.0152 | 0.84558
| Rel.Std.Dev. |  | 2.923% | 1.82% | 3.795% | 1.96% | 2.979% | 3.13% | 2.378% | 0% | 4.221% |  | 6.855% | 2.425% | 7.881% | 7.787% | 19.75% | 6.606% | 6.499% | 1.661% | 8.281% | 6.555% | 6.411% | 3.162% | 12.89% | 4.622% | 3.72% | 5.212% | 5.261% | 5.779% | 4.711%
| PDM3 |  | 0.2731% | 4.721% | 12.57% | 18.25% | -1.619% | -0.6728% | -2.565% | 13.52% | 1.219% |  | 0.2731% | -4.457% | -15.81% | 9.733% | 134.6% | 4.057% | -19.31% | 30.45% | 28.37% | 1.219% | 22.03% | -7.295% | -10.13% | 0.2731% | -3.511% | -11.08% | -3.511% | -0.2944% | 1.881%
| Z-Score (Absolute) |  | 0 | 0.65 | 1.81 | 2.64 | 0.28 | 0.14 | 0.42 | 1.95 | 0.14 | NA | 0 | 0.7 | 2.36 | 1.39 | 19.74 | 0.56 | 2.88 | 4.43 | 3.55 | 0.14 | 2.78 | 1.11 | 2.22 | 0 | 0.56 | 1.67 | 0.56 | 0.41 | 0.24
| Analytical results for Cs in OREAS 905 (Certified Value 1.32 ppm) | [31 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 17 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [10 empty] | 2
| 1 | 1 | 1.35 | 1.6245 | 1.49 | 1.29 | 1.13 | 1.41 | 1.73 | 1.2 | 1.44 | 1.29 | 1.17 | 1.45 | 1.33 | 1.6 | 0.98 | 1.32 | 1.44 | 0.58 | 1.22 | 1.46 | [10 empty] | 1
| 1 | 2 | 1.23 | 1.5911 | 1.39 | 1.33 | 1.09 | 1.4 | 1.81 | 1.2 | 1.56 | 1.24 | 1.23 | 1.46 | 1.26 | 1.6 | 1.02 | 1.28 | 1.33 | 0.58 | 1.15 | 1.42 | [10 empty] | 0
| 1 | 3 | 1.33 | 1.5119 | 1.44 | 1.31 | 1.12 | 1.37 | 1.9 | 1.1 | 1.39 | 1.34 | 1.22 | 1.5 | 1.29 | 1.6 | 0.86 | 1.26 | 1.42 | 0.59 | 1.15 | 1.41 | [10 empty] | 16
| 1 | 4 | 1.3 | 1.4583 | 1.47 | 1.31 | 1.13 | 1.39 | 1.86 | 1.1 | 1.37 | 1.15 | 1.22 | 1.52 | 1.27 | 1.6 | 1.06 | 1.2 | 1.49 | 0.63 | 1.21 | 1.44 | [10 empty] | 1.3175
| 1 | 5 | 1.34 | 1.554 | 1.44 | 1.37 | 1.1 | 1.34 | 1.81 | 1.1 | 1.44 | 1.23 | 1.22 | 1.19 | 1.26 | 1.6 | 1.06 | 1.24 | 1.27 | 0.57 | 1.16 | 1.41 | [10 empty] | 130
| 1 | 6 | 1.38 | 1.5659 | 1.45 | 1.26 | 1.1 | 1.38 | 1.73 | 1.1 | 1.36 | 1.37 | 1.12 | 1.49 | 1.29 | 1.6 | 0.98 | 1.23 | 1.08 | 0.59 | 1.15 | 1.39
| Mean |  | 1.3217 | 1.5509 | 1.4467 | 1.3117 | 1.1117 | 1.3817 | 1.8067 | 1.1333 | 1.4267 | 1.27 | 1.1967 | 1.435 | 1.2833 | 1.6 | 0.99333 | 1.255 | 1.3383 | 0.59 | 1.1733 | 1.4217
| Median |  | 1.335 | 1.5599 | 1.445 | 1.31 | 1.11 | 1.385 | 1.81 | 1.1 | 1.415 | 1.265 | 1.22 | 1.475 | 1.28 | 1.6 | 1 | 1.25 | 1.375 | 0.585 | 1.155 | 1.415
| Std Dev. |  | 0.051929 | 0.058931 | 0.033862 | 0.037103 | 0.017224 | 0.024833 | 0.068313 | 0.05164 | 0.073666 | 0.08025 | 0.043205 | 0.12276 | 0.026583 | 2.4324e-16 | 0.074476 | 0.041833 | 0.14932 | 0.020976 | 0.03266 | 0.024833
| Rel.Std.Dev. |  | 3.929% | 3.8% | 2.341% | 2.829% | 1.549% | 1.797% | 3.781% | 4.556% | 5.163% | 6.319% | 3.61% | 8.555% | 2.071% | 1.52e-14% | 7.498% | 3.333% | 11.16% | 3.555% | 2.784% | 1.747%
| PDM3 |  | 0.3151% | 17.72% | 9.803% | -0.4439% | -15.62% | 4.869% | 37.13% | -13.98% | 8.285% | -3.606% | -9.172% | 8.917% | -2.594% | 21.44% | -24.61% | -4.745% | 1.58% | -55.22% | -10.94% | 7.905%
| Z-Score (Absolute) |  | 0.03 | 1.32 | 0.73 | 0.03 | 1.15 | 0.37 | 2.75 | 1.03 | 0.62 | 0.26 | 0.59 | 0.94 | 0.19 | 1.59 | 1.82 | 0.35 | 0.12 | 4.13 | 0.81 | 0.59
| Analytical results for Cu in OREAS 905 (Certified Value 1562 ppm) | [31 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | ppm
|  |  | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*AAS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | 0
|  |  | 0.4g | 01g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.5g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g | 0
| 1 | 1 | 1595 | 1600 | 1528.716 | 1577 | 1490 | 1575 | 1540 | 1530 | 1620 | 1380 | 1600 | 1600 | 1577 | 1684 | 1530 | 1462 | 1567 | 1558 | 1538 | 1455 | 1461 | 1680 | 1590 | 1580 | 1530 | 1540 | 1660 | 1560 | 1590 | 1571.66 | 1
| 1 | 2 | 1560 | 1510 | 1512.2724 | 1557 | 1490 | 1590 | 1570 | 1490 | 1630 | 1430 | 1600 | 1570 | 1592 | 1604 | 1550 | 1480 | 1521 | 1487 | 1536 | 1565 | 1513 | 1520 | 1617 | 1550 | 1530 | 1550 | 1660 | 1525 | 1510 | 1555.28 | 0
| 1 | 3 | 1595 | 1510 | 1525.863 | 1557 | 1490 | 1590 | 1540 | 1480 | 1600 | 1390 | 1600 | 1570 | 1590 | 1627 | 1600 | 1445 | 1609 | 1400 | 1547 | 1570 | 1547 | 1560 | 1612 | 1550 | 1550 | 1540 | 1650 | 1575 | 1530 | 1541.2 | 16
| 1 | 4 | 1590 | 1540 | 1514.477 | 1569 | 1500 | 1565 | 1560 | 1490 | 1660 | 1450 | 1600 | 1560 | 1562 | 1634 | 1590 | 1443 | 1555 | 1427 | 1559 | 1485 | 1545 | 1580 | 1595 | 1570 | 1520 | 1540 | 1660 | 1595 | 1520 | 1569.57 | 1561.789
| 1 | 5 | 1605 | 1580 | 1516.4792 | 1564 | 1490 | 1575 | 1520 | 1510 | 1620 | 1460 | 1600 | 1600 | 1626 | 1617 | 1590 | 1505 | 1544 | 1388 | 1568 | 1550 | 1546 | 1560 | 1634 | 1570 | 1520 | 1530 | 1680 | 1600 | 1550 | 1559.06 | 131
| 1 | 6 | 1620 | 1550 | 1523.8614 | 1564 | 1510 | 1570 | 1560 | 1510 | 1630 | 1410 | 1600 | 1570 | 1607 | 1641 | 1540 | 1533 | 1572 | 1393 | 1540 | 1575 | 1498 | 1580 | 1648 | 1560 | 1530 | 1560 | 1650 | 1600 | 1520 | 1612.12
| Mean |  | 1594.1667 | 1548.3333 | 1520.2782 | 1564.6667 | 1495 | 1577.5 | 1548.3333 | 1501.6667 | 1626.6667 | 1420 | 1600 | 1578.3333 | 1592.3333 | 1634.5 | 1566.6667 | 1478 | 1561.3333 | 1442.1667 | 1548 | 1533.3333 | 1518.3333 | 1580 | 1616 | 1563.3333 | 1530 | 1543.3333 | 1660 | 1575.8333 | 1536.6667 | 1568.1483
| Median |  | 1595 | 1545 | 1520.1703 | 1564 | 1490 | 1575 | 1550 | 1500 | 1625 | 1420 | 1600 | 1570 | 1591 | 1630.5 | 1570 | 1471 | 1561 | 1413.5 | 1543.5 | 1557.5 | 1529 | 1570 | 1614.5 | 1565 | 1530 | 1540 | 1660 | 1585 | 1525 | 1564.315
| Std Dev. |  | 19.8536 | 36.5605 | 6.744 | 7.607 | 8.3666 | 10.3682 | 18.3485 | 18.3485 | 19.6638 | 32.249 | 0 | 17.224 | 22.4024 | 27.5155 | 30.1109 | 35.5753 | 29.6153 | 67.5083 | 12.8841 | 50.6623 | 34.7198 | 53.6656 | 22.28 | 12.1106 | 10.9545 | 10.328 | 10.9545 | 29.5663 | 29.4392 | 24.1755
| Rel.Std.Dev. |  | 1.245% | 2.361% | 0.4436% | 0.4862% | 0.5596% | 0.6573% | 1.185% | 1.222% | 1.209% | 2.271% | 0% | 1.091% | 1.407% | 1.683% | 1.922% | 2.407% | 1.897% | 4.681% | 0.8323% | 3.304% | 2.287% | 3.397% | 1.379% | 0.7747% | 0.716% | 0.6692% | 0.6599% | 1.876% | 1.916% | 1.542%
| PDM3 |  | 2.073% | -0.8616% | -2.658% | 0.1843% | -4.276% | 1.006% | -0.8616% | -3.85% | 4.154% | -9.079% | 2.447% | 1.059% | 1.956% | 4.656% | 0.3123% | -5.365% | -0.02918% | -7.659% | -0.8829% | -1.822% | -2.782% | 1.166% | 3.471% | 0.09888% | -2.035% | -1.182% | 6.288% | 0.8992% | -1.609% | 0.4072%
| Z-Score (Absolute) |  | 0.72 | 0.27 | 0.87 | 0.09 | 1.42 | 0.36 | 0.27 | 1.28 | 1.43 | 3.04 | 0.85 | 0.38 | 0.69 | 1.6 | 0.13 | 1.79 | 0.01 | 2.56 | 0.27 | 0.59 | 0.92 | 0.01 | 1.2 | 0.06 | 0.66 | 0.38 | 2.15 | 0.33 | 0.75 | 0.16
| Analytical results for Dy in OREAS 905 (Certified Value 1.73 ppm) | [31 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 03; 02: 05; 03: 09; 04: 12; 05: 13; 06: 17; 07: 22; 08: 26; 09: 30 | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS; 09: AR*MS | 2
|  |  | 01: 20g; 02: 0.4g; 03: 40g; 04: 0.5g; 05: 25g; 06: 0.5g; 07: 25g; 08: 40g; 09: 0.4g | 2
| 1 | 1 | 01: 2.2207; 02: 1.81; 03: 2.13; 04: 1.5; 05: 1.62; 06: 1.76; 07: 1.2; 08: 1.72; 09: 1.39 | 1
| 1 | 2 | 01: 2.0933; 02: 1.85; 03: 2.23; 04: 1.3; 05: 1.66; 06: 1.73; 07: 1.3; 08: 2.04; 09: 1.37 | 0
| 1 | 3 | 01: 2.0639; 02: 1.85; 03: 2.18; 04: 1.4; 05: 1.72; 06: 1.81; 07: 1.3; 08: 1.9; 09: 1.39 | 16
| 1 | 4 | 01: 2.0488; 02: 1.84; 03: 2.23; 04: 1.4; 05: 1.71; 06: 1.71; 07: 1.2; 08: 2.11; 09: 1.42 | 1.7302
| 1 | 5 | 01: 2.1021; 02: 1.9; 03: 2.26; 04: 1.4; 05: 1.69; 06: 1.74; 07: 1.3; 08: 1.85; 09: 1.39 | 132
| 1 | 6 | 01: 2.2405; 02: 1.8; 03: 2.19; 04: 1.4; 05: 1.51; 06: 1.75; 07: 1.3; 08: 2.05; 09: 1.35
| Mean |  | 01: 2.1282; 02: 1.8417; 03: 2.2033; 04: 1.4; 05: 1.6517; 06: 1.75; 07: 1.2667; 08: 1.945; 09: 1.385
| Median |  | 01: 2.0977; 02: 1.845; 03: 2.21; 04: 1.4; 05: 1.675; 06: 1.745; 07: 1.3; 08: 1.97; 09: 1.39
| Std Dev. |  | 01: 0.081885; 02: 0.035449; 03: 0.046332; 04: 0.063246; 05: 0.078337; 06: 0.034059; 07: 0.05164; 08: 0.14761; 09: 0.023452
| Rel.Std.Dev. |  | 01: 3.848%; 02: 1.925%; 03: 2.103%; 04: 4.518%; 05: 4.743%; 06: 1.946%; 07: 4.077%; 08: 7.589%; 09: 1.693%
| PDM3 |  | 01: 23.01%; 02: 6.444%; 03: 27.35%; 04: -19.08%; 05: -4.537%; 06: 1.146%; 07: -26.79%; 08: 12.42%; 09: -19.95%
| Z-Score (Absolute) |  | 01: 0.73; 02: 0.18; 03: 0.87; 04: 0.67; 05: 0.19; 06: 0; 07: 0.93; 08: 0.38; 09: 0.7
| Analytical results for Er in OREAS 905 (Indicative Value 0.44 ppm) | [31 empty] | Indicative
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 03; 02: 05; 03: 09; 04: 12; 05: 13; 06: 17; 07: 22; 08: 26; 09: 30 | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS; 09: AR*MS | 2
|  |  | 01: 20g; 02: 0.4g; 03: 40g; 04: 0.5g; 05: 25g; 06: 0.5g; 07: 25g; 08: 40g; 09: 0.4g | 2
| 1 | 1 | 01: 0.589; 02: 0.54; 03: 0.42; 04: 0.4; 05: 0.42; 06: 0.47; 07: 0.3; 08: 0.38; 09: 1.8 | 1
| 1 | 2 | 01: 0.59197; 02: 0.55; 03: 0.41; 04: 0.4; 05: 0.43; 06: 0.46; 07: 0.3; 08: 0.47; 09: 1.77 | 13
| 1 | 3 | 01: 0.54304; 02: 0.56; 03: 0.45; 04: 0.4; 05: 0.44; 06: 0.47; 07: 0.3; 08: 0.41; 09: 1.77 | 16
| 1 | 4 | 01: 0.53561; 02: 0.55; 03: 0.43; 04: 0.3; 05: 0.43; 06: 0.43; 07: 0.3; 08: 0.45; 09: 1.8 | 0.44082
| 1 | 5 | 01: 0.56969; 02: 0.58; 03: 0.42; 04: 0.4; 05: 0.43; 06: 0.45; 07: 0.3; 08: 0.4; 09: 1.84 | 19
| 1 | 6 | 01: 0.5902; 02: 0.54; 03: 0.42; 04: 0.3; 05: 0.39; 06: 0.46; 07: 0.3; 08: 0.44; 09: 1.77
| Mean |  | 01: 0.56992; 02: 0.55333; 03: 0.425; 04: 0.36667; 05: 0.42333; 06: 0.45667; 07: 0.3; 08: 0.425; 09: 1.7917
| Median |  | 01: 0.57934; 02: 0.55; 03: 0.42; 04: 0.4; 05: 0.43; 06: 0.46; 07: 0.3; 08: 0.425; 09: 1.785
| Std Dev. |  | 01: 0.025146; 02: 0.015055; 03: 0.013784; 04: 0.05164; 05: 0.017512; 06: 0.015055; 07: 0; 08: 0.033912; 09: 0.027869
| Rel.Std.Dev. |  | 01: 4.412%; 02: 2.721%; 03: 3.243%; 04: 14.08%; 05: 4.137%; 06: 3.297%; 07: 0%; 08: 7.979%; 09: 1.555%
| PDM3 |  | 01: 29.28%; 02: 25.52%; 03: -3.589%; 04: -16.82%; 05: -3.968%; 06: 3.594%; 07: -31.95%; 08: -3.589%; 09: 306.4%
| Z-Score (Absolute) |  | 01: 1.49; 02: 1.31; 03: 0.05; 04: 0.67; 05: 0; 06: 0.28; 07: 1.38; 08: 0.05; 09: 14.5
| Analytical results for Eu in OREAS 905 (Certified Value 0.92 ppm) | [31 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | 1
| No | No | 01: 03; 02: 05; 03: 09; 04: 12; 05: 13; 06: 17; 07: 22; 08: 26; 09: 30 | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS; 09: AR*MS | 2
|  |  | 01: 20g; 02: 0.4g; 03: 40g; 04: 0.5g; 05: 25g; 06: 0.5g; 07: 25g; 08: 40g; 09: 0.4g | 3
| 1 | 1 | 01: 1.058; 02: 0.92; 03: 0.926; 04: 0.9; 05: 0.8; 06: 0.91; 07: 0.5; 08: 0.754; 09: 1.05 | 1
| 1 | 2 | 01: 1.0717; 02: 0.94; 03: 0.926; 04: 0.9; 05: 0.8; 06: 0.9; 07: 0.6; 08: 0.834; 09: 1.02 | 0
| 1 | 3 | 01: 1.0611; 02: 0.95; 03: 0.928; 04: 0.9; 05: 0.85; 06: 0.94; 07: 0.6; 08: 0.736; 09: 1.01 | 16
| 1 | 4 | 01: 1.04; 02: 0.96; 03: 0.936; 04: 0.8; 05: 0.84; 06: 0.89; 07: 0.6; 08: 0.84; 09: 1.04 | 0.91738
| 1 | 5 | 01: 1.0476; 02: 0.97; 03: 0.918; 04: 0.9; 05: 0.82; 06: 0.9; 07: 0.6; 08: 0.77; 09: 1.03 | 133
| 1 | 6 | 01: 1.0618; 02: 0.91; 03: 0.906; 04: 0.9; 05: 0.77; 06: 0.9; 07: 0.6; 08: 0.83; 09: 1.01
| Mean |  | 01: 1.0567; 02: 0.94167; 03: 0.92333; 04: 0.88333; 05: 0.81333; 06: 0.90667; 07: 0.58333; 08: 0.794; 09: 1.0267
| Median |  | 01: 1.0595; 02: 0.945; 03: 0.926; 04: 0.9; 05: 0.81; 06: 0.9; 07: 0.6; 08: 0.8; 09: 1.025
| Std Dev. |  | 01: 0.011253; 02: 0.023166; 03: 0.01025; 04: 0.040825; 05: 0.029439; 06: 0.017512; 07: 0.040825; 08: 0.045939; 09: 0.01633
| Rel.Std.Dev. |  | 01: 1.065%; 02: 2.46%; 03: 1.11%; 04: 4.622%; 05: 3.62%; 06: 1.931%; 07: 6.999%; 08: 5.786%; 09: 1.591%
| PDM3 |  | 01: 15.18%; 02: 2.648%; 03: 0.6493%; 04: -3.711%; 05: -11.34%; 06: -1.167%; 07: -36.41%; 08: -13.45%; 09: 11.91%
| Z-Score (Absolute) |  | 01: 1.22; 02: 0.32; 03: 0.18; 04: 0.13; 05: 0.67; 06: 0; 07: 2.46; 08: 0.82; 09: 0.99
| Analytical results for Fe in OREAS 905 (Certified Value 3.5 wt.%) | [31 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 3.4 | 3.5773 | 3.35 | 3.68 | 3.5 | 3.43 | 3.92 | 4.04 | 3.33 | 3.75 | 3.58 | 3.3 | 3.68 | 3.6 | 3.5 | 3.46 | 3.53 | 3.3555 | 3.21 | 2.08 | 3.44 | 3.78 | 3.36 | 3.35 | 3.82 | 3.48 | 3.56 | 3.25 |  |  | 1
| 1 | 2 | 3.32 | 3.52 | 3.3 | 3.68 | 3.5 | 3.52 | 4.03 | 4.04 | 3.52 | 3.75 | 3.53 | 3.38 | 3.49 | 3.4 | 3.57 | 3.35 | 3.485 | 3.6205 | 3.4 | 2.14 | 3.52 | 3.81 | 3.34 | 3.37 | 3.79 | 3.36 | 3.37 | 3.21 |  |  | 0
| 1 | 3 | 3.39 | 3.5457 | 3.3 | 3.52 | 3.55 | 3.45 | 4.04 | 4.08 | 3.42 | 3.79 | 3.53 | 3.35 | 3.61 | 3.51 | 3.42 | 3.55 | 3.435 | 3.63 | 3.37 | 2.11 | 3.51 | 3.8 | 3.36 | 3.36 | 3.84 | 3.43 | 3.43 | 3.24 |  |  | 16
| 1 | 4 | 3.4 | 3.5081 | 3.31 | 3.57 | 3.46 | 3.49 | 4.07 | 4.17 | 3.52 | 3.7 | 3.6 | 3.29 | 3.63 | 3.49 | 3.48 | 3.44 | 3.55 | 3.495 | 3.4 | 2.24 | 3.44 | 3.78 | 3.32 | 3.35 | 3.83 | 3.52 | 3.41 | 3.29 |  |  | 3.4952
| 1 | 5 | 3.41 | 3.5428 | 3.26 | 3.56 | 3.46 | 3.39 | 4.07 | 4.07 | 3.47 | 3.7 | 3.57 | 3.44 | 3.57 | 3.52 | 3.63 | 3.39 | 3.63 | 3.57 | 3.41 | 2.2 | 3.52 | 3.94 | 3.31 | 3.3 | 3.87 | 3.5 | 3.48 | 3.26 |  |  | 134
| 1 | 6 | 3.4 | 3.5469 | 3.3 | 3.63 | 3.48 | 3.45 | 4 | 4.1 | 3.44 | 3.8 | 3.56 | 3.28 | 3.61 | 3.39 | 3.66 | 3.46 | 3.505 | 3.68 | 3.33 | 2.16 | 3.49 | 3.8 | 3.32 | 3.35 | 3.83 | 3.47 | 3.42 | 3.35
| Mean |  | 3.3867 | 3.5401 | 3.3033 | 3.6067 | 3.4917 | 3.455 | 4.0217 | 4.0833 | 3.45 | 3.7483 | 3.5617 | 3.34 | 3.5983 | 3.485 | 3.5433 | 3.4417 | 3.5225 | 3.5585 | 3.3533 | 2.155 | 3.4867 | 3.8183 | 3.335 | 3.3467 | 3.83 | 3.46 | 3.445 | 3.2667
| Median |  | 3.4 | 3.5443 | 3.3 | 3.6 | 3.49 | 3.45 | 4.035 | 4.075 | 3.455 | 3.75 | 3.565 | 3.325 | 3.61 | 3.5 | 3.535 | 3.45 | 3.5175 | 3.5953 | 3.385 | 2.15 | 3.5 | 3.8 | 3.33 | 3.35 | 3.83 | 3.475 | 3.425 | 3.255
| Std Dev. |  | 0.033267 | 0.024069 | 0.028752 | 0.066833 | 0.033714 | 0.045497 | 0.056362 | 0.048442 | 0.071554 | 0.042622 | 0.027869 | 0.06229 | 0.064005 | 0.079183 | 0.092664 | 0.068532 | 0.065936 | 0.11755 | 0.07607 | 0.058566 | 0.037771 | 0.0608 | 0.021679 | 0.024221 | 0.026077 | 0.057619 | 0.066558 | 0.048442
| Rel.Std.Dev. |  | 0.9823% | 0.6799% | 0.8704% | 1.853% | 0.9656% | 1.317% | 1.401% | 1.186% | 2.074% | 1.137% | 0.7825% | 1.865% | 1.779% | 2.272% | 2.615% | 1.991% | 1.872% | 3.303% | 2.268% | 2.718% | 1.083% | 1.592% | 0.6501% | 0.7237% | 0.6809% | 1.665% | 1.932% | 1.483%
| PDM3 |  | -3.104% | 1.287% | -5.488% | 3.191% | -0.09972% | -1.149% | 15.06% | 16.83% | -1.292% | 7.244% | 1.903% | -4.439% | 2.952% | -0.2905% | 1.379% | -1.53% | 0.7825% | 1.812% | -4.058% | -38.34% | -0.2428% | 9.247% | -4.582% | -4.248% | 9.58% | -1.006% | -1.435% | -6.537%
| Z-Score (Absolute) |  | 0.66 | 0.36 | 1.21 | 0.8 | 0.04 | 0.2 | 3.56 | 3.97 | 0.24 | 1.74 | 0.5 | 0.97 | 0.75 | 0.01 | 0.38 | 0.29 | 0.24 | 0.48 | 0.69 | 8.84 | 0.01 | 2.05 | 1 | 0.92 | 2.29 | 0.17 | 0.27 | 1.46
| Analytical results for Ga in OREAS 905 (Certified Value 6.43 ppm) | [31 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 6.97 | 7.3718 | 6.6 | 6.2 | 5.95 | 6.74 | 5.67 | 8.4 | 7.3 | 7 | 6.58 | 6.4 | 6.8 | 15 | 6.34 | 7 | 6.1 | 3.1 | < 5 | 7 | 7 | 6.2 | 6.75 | 5.46 | < 0.2 | 5.96 | 7.2 | 5.3 |  |  | 1
| 1 | 2 | 6.36 | 7.2826 | 6.2 | 6.3 | 5.87 | 6.72 | 6.05 | 7.8 | 7.5 | 7 | 6.34 | 6.3 | 6.7 | 14 | 6.12 | 7 | 5.9 | 3.3 | < 5 | 10 | 10 | 6.4 | 6.53 | 5.38 | < 0.2 | 5.91 | 6.8 | 5.3 |  |  | 0
| 1 | 3 | 6.93 | 7.2793 | 6.1 | 6.3 | 5.97 | 6.5 | 5.79 | 8.3 | 6.9 | 6 | 7.08 | 6.6 | 7 | 12 | 6.27 | 7 | 5.8 | 3.2 | < 5 | 10 | < 5 | 5.4 | 6.48 | 5.4 | < 0.2 | 6.27 | 7 | 5.2 |  |  | 16
| 1 | 4 | 6.79 | 7.2895 | 6.4 | 6.3 | 6.1 | 6.67 | 6.17 | 7.9 | 7.1 | 7 | 6.18 | 6.6 | 7 | 13 | 6.14 | 7 | 6.1 | 3.1 | < 5 | 11 | 7 | 7 | 5.99 | 5.84 | < 0.2 | 6.16 | 6.7 | 5.5 |  |  | 6.4321
| 1 | 5 | 6.77 | 7.3447 | 6.1 | 6.4 | 5.97 | 6.46 | 5.95 | 8.3 | 7.2 | 6 | 7.16 | 6.8 | 6.5 | 12 | 6.15 | 7 | 5.9 | 3.2 | 5 | 10 | 5 | 6.4 | 6.21 | 5.45 | < 0.2 | 6.18 | 6.7 | 5.5 |  |  | 135
| 1 | 6 | 6.98 | 7.3546 | 6.4 | 6.4 | 5.88 | 6.58 | 5.58 | 7.9 | 6.9 | 7 | 7.28 | 6.2 | 7 | 15 | 6.18 | 7 | 6 | 3.5 | < 5 | 11 | < 5 | 6.6 | 6.28 | 5.47 | < 0.2 | 6.22 | 6.7 | 5.5
| Mean |  | 6.8 | 7.3204 | 6.3 | 6.3167 | 5.9567 | 6.6117 | 5.8683 | 8.1 | 7.15 | 6.6667 | 6.77 | 6.4833 | 6.8333 | 13.5 | 6.2 | 7 | 5.9667 | 3.2333 | 5 | 9.8333 | 7.25 | 6.3333 | 6.3733 | 5.5 |  | 6.1167 | 6.85 | 5.3833
| Median |  | 6.86 | 7.3171 | 6.3 | 6.3 | 5.96 | 6.625 | 5.87 | 8.1 | 7.15 | 7 | 6.83 | 6.5 | 6.9 | 13.5 | 6.165 | 7 | 5.95 | 3.2 | 5 | 10 | 7 | 6.4 | 6.38 | 5.455 |  | 6.17 | 6.75 | 5.4
| Std Dev. |  | 0.23358 | 0.041167 | 0.2 | 0.075277 | 0.082865 | 0.11669 | 0.22772 | 0.26077 | 0.23452 | 0.5164 | 0.4642 | 0.22286 | 0.20656 | 1.3784 | 0.086487 | 0 | 0.12111 | 0.15055 |  | 1.472 | 2.0616 | 0.53166 | 0.26838 | 0.17029 |  | 0.14652 | 0.20736 | 0.13292
| Rel.Std.Dev. |  | 3.435% | 0.5624% | 3.175% | 1.192% | 1.391% | 1.765% | 3.88% | 3.219% | 3.28% | 7.746% | 6.857% | 3.437% | 3.023% | 10.21% | 1.395% | 0% | 2.03% | 4.656% |  | 14.97% | 28.44% | 8.395% | 4.211% | 3.096% |  | 2.395% | 3.027% | 2.469%
| PDM3 |  | 5.72% | 13.81% | -2.054% | -1.795% | -7.392% | 2.792% | -8.765% | 25.93% | 11.16% | 3.647% | 5.253% | 0.7965% | 6.238% | 109.9% | -3.608% | 8.829% | -7.236% | -49.73% | -22.26% | 52.88% | 12.72% | -1.536% | -0.9136% | -14.49% |  | -4.904% | 6.497% | -16.31%
| Z-Score (Absolute) |  | 0.67 | 1.39 | 0.31 | 0.28 | 0.88 | 0.21 | 1.02 | 2.69 | 1.11 | 0.31 | 0.48 | 0 | 0.58 | 11.69 | 0.47 | 0.86 | 0.86 | 5.42 | 2.47 | 6.53 | 0.25 | 0.06 | 0.18 | 1.75 | NA | 0.61 | 0.49 | 1.83
| Analytical results for Gd in OREAS 905 (Certified Value 3.59 ppm) | [31 empty] | Certified
Gd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 12; 06: 13; 07: 17; 08: 22; 09: 26; 10: 30 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*OES/MS; 09: AR*MS; 10: AR*MS |  |  | 2
|  |  | 01: 20g; 03: 0.4g; 04: 40g; 05: 0.5g; 06: 25g; 07: 0.5g; 08: 25g; 09: 40g; 10: 0.4g |  |  | 3
| 1 | 1 | 01: 4.2029; 03: 3.77; 04: 3.32; 05: 3.5; 06: 2.99; 07: 3.95; 08: 3.3; 09: 3.04; 10: 11.9 |  |  | 1
| 1 | 2 | 01: 4.0499; 03: 3.92; 04: 3.27; 05: 3.5; 06: 3.14; 07: 3.87; 08: 3.6; 09: 3.55; 10: 11.5 |  |  | 0
| 1 | 3 | 01: 4.1509; 03: 3.89; 04: 3.41; 05: 3.5; 06: 3.17; 07: 3.93; 08: 3.6; 09: 3.26; 10: 11.5 |  |  | 16
| 1 | 4 | 01: 3.9018; 03: 3.9; 04: 3.33; 05: 3.5; 06: 3.19; 07: 3.81; 08: 3.4; 09: 3.77; 10: 11.6 |  |  | 3.5854
| 1 | 5 | 01: 4.0006; 03: 4; 04: 3.35; 05: 3.6; 06: 3.14; 07: 3.8; 08: 3.6; 09: 3.11; 10: 11.6 |  |  | 136
| 1 | 6 | 01: 4.3038; 03: 3.76; 04: 3.21; 05: 3.5; 06: 2.94; 07: 3.88; 08: 3.7; 09: 3.52; 10: 11.6
| Mean |  | 01: 4.1016; 03: 3.8733; 04: 3.315; 05: 3.5167; 06: 3.095; 07: 3.8733; 08: 3.5333; 09: 3.375; 10: 11.6167
| Median |  | 01: 4.1004; 03: 3.895; 04: 3.325; 05: 3.5; 06: 3.14; 07: 3.875; 08: 3.6; 09: 3.39; 10: 11.6
| Std Dev. |  | 01: 0.14586; 03: 0.092448; 04: 0.068629; 05: 0.040825; 06: 0.10368; 07: 0.060882; 08: 0.15055; 09: 0.28402; 10: 0.1472
| Rel.Std.Dev. |  | 01: 3.556%; 03: 2.387%; 04: 2.07%; 05: 1.161%; 06: 3.35%; 07: 1.572%; 08: 4.261%; 09: 8.416%; 10: 1.267%
| PDM3 |  | 01: 14.4%; 03: 8.03%; 04: -7.542%; 05: -1.917%; 06: -13.68%; 07: 8.03%; 08: -1.453%; 09: -5.869%; 10: 224%
| Z-Score (Absolute) |  | 01: 1.13; 03: 0.67; 04: 0.43; 05: 0.03; 06: 0.87; 07: 0.67; 08: 0; 09: 0.31; 10: 16.03
| Analytical results for Ge in OREAS 905 (Certified Value < 1 ppm) | [31 empty] | Certified
Ge | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 15: Lab; 16: Lab |  |  | 1
| No | No | 01: 01; 03: 04; 04: 06; 05: 07; 06: 10; 07: 11; 08: 12; 09: 15; 10: 22; 11: 24; 12: 28; 13: 29; 15: 31; 16: 32 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*OES/MS; 04: AR*OES/MS; 05: AR*OES/MS; 06: AR*OES/MS; 07: AR*OES; 08: AR*MS; 09: AR*OES/MS; 10: AR*OES/MS; 11: AR*OES/MS; 12: AR*OES/MS; 13: AR*OES/MS; 15: AR*OES/MS; 16: AR*OES/MS |  |  | 2
|  |  | 01: 0.4g; 03: 0.25g; 04: 0.5g; 05: 0.5g; 06: 0.25g; 07: 0.15g; 08: 0.5g; 09: 0.25g; 10: 25g; 11: 0.4g; 12: 0.5g; 13: 0.5g; 15: 0.4g; 16: 0.25g |  |  | 2
| 1 | 1 | 01: 0.11; 03: 0.1; 04: 0.09; 05: 0.09; 06: < 0.1; 07: < 10; 08: 0.4; 09: 0.1; 10: < 0.2; 11: 2.3; 12: 0.1; 13: 0.15; 15: 0.1; 16: < 0.1 |  |  | 1
| 1 | 2 | 01: 0.1; 03: 0.2; 04: 0.1; 05: 0.09; 06: < 0.1; 07: < 10; 08: 0.5; 09: 0.1; 10: < 0.2; 11: 2.3; 12: 0.08; 13: 0.14; 15: 0.11; 16: < 0.1 |  |  | 4
| 1 | 3 | 01: 0.11; 03: 0.2; 04: 0.11; 05: 0.09; 06: < 0.1; 07: < 10; 08: 0.5; 09: 0.1; 10: < 0.2; 11: 2.3; 12: 0.1; 13: 0.14; 15: 0.12; 16: < 0.1 |  |  | 16
| 1 | 4 | 01: 0.11; 03: 0.2; 04: 0.1; 05: 0.08; 06: < 0.1; 07: < 10; 08: 0.5; 09: 0.1; 10: < 0.2; 11: 2.3; 12: 0.08; 13: 0.15; 15: 0.12; 16: < 0.1 |  |  | < 1
| 1 | 5 | 01: 0.11; 03: 0.2; 04: 0.09; 05: 0.08; 06: < 0.1; 07: < 10; 08: 0.5; 09: 0.1; 10: < 0.2; 11: 2.3; 12: 0.08; 13: 0.13; 15: 0.11; 16: < 0.1 |  |  | 137
| 1 | 6 | 01: 0.12; 03: 0.2; 04: 0.1; 05: 0.1; 06: < 0.1; 07: < 10; 08: 0.5; 09: 0.1; 10: < 0.2; 11: 2.3; 12: 0.08; 13: 0.15; 15: 0.11; 16: < 0.1
| Mean |  | 01: 0.11; 03: 0.18333; 04: 0.098333; 05: 0.088333; 08: 0.48333; 09: 0.1; 11: 2.3; 12: 0.086667; 13: 0.14333; 15: 0.11167
| Median |  | 01: 0.11; 03: 0.2; 04: 0.1; 05: 0.09; 08: 0.5; 09: 0.1; 11: 2.3; 12: 0.08; 13: 0.145; 15: 0.11
| Std Dev. |  | 01: 0.0063246; 03: 0.040825; 04: 0.0075277; 05: 0.0075277; 08: 0.040825; 09: 1.5202e-17; 11: 4.8648e-16; 12: 0.010328; 13: 0.008165; 15: 0.0075277
| Rel.Std.Dev. |  | 01: 5.75%; 03: 22.27%; 04: 7.655%; 05: 8.522%; 08: 8.447%; 09: 1.52e-14%; 11: 2.115e-14%; 12: 11.92%; 13: 5.696%; 15: 6.741%
PDM3
| Z-Score (Absolute) |  | 01: 0.02; 03: 2.1; 04: 0.36; 05: 0.65; 06: NA; 07: NA; 08: 10.76; 09: 0.31; 10: NA; 11: 63.26; 12: 0.7; 13: 0.94; 15: 0.02; 16: NA
| Analytical results for Hf in OREAS 905 (Certified Value 1.22 ppm) | [31 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 15 | 17 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [10 empty] | 2
| 1 | 1 | 1.23 | 1.7263 | 1.5 | 1.26 | 1.25 | 1.29 | 1.3 | 0.82 | 1 | 0.16 | 1.35 | 1.25 | 0.26 | 1.1 | 0.6 | 1.25 | 1.04 | 1.02 | 1.24 | 1.41 | [10 empty] | 1
| 1 | 2 | 1.14 | 1.758 | 1.38 | 1.34 | 1.21 | 1.29 | 1.31 | 0.92 | 0.7 | 0.18 | 1.39 | 1.22 | 0.27 | 1.6 | < 0.2 | 1.18 | 1.04 | 1.06 | 1.16 | 1.34 | [10 empty] | 0
| 1 | 3 | 1.28 | 1.747 | 1.43 | 1.21 | 1.25 | 1.32 | 1.26 | 0.91 | 0.9 | 0.15 | 1.44 | 1.2 | 0.23 | 1.6 | < 0.2 | 1.23 | 1.13 | 1.02 | 1.15 | 1.31 | [10 empty] | 16
| 1 | 4 | 1.16 | 1.709 | 1.42 | 1.22 | 1.26 | 1.32 | 1.31 | 0.85 | 1.1 | 0.16 | 1.43 | 1.16 | 0.35 | 1.5 | < 0.2 | 1.22 | 1.12 | 1.09 | 1.18 | 1.39 | [10 empty] | 1.2219
| 1 | 5 | 1.18 | 1.7635 | 1.46 | 1.26 | 1.24 | 1.3 | 1.25 | 0.94 | 1.1 | 0.12 | 1.32 | 1.18 | 0.13 | 1.5 | 0.8 | 1.16 | 0.97 | 1.02 | 1.17 | 1.32 | [10 empty] | 138
| 1 | 6 | 1.24 | 1.7718 | 1.46 | 1.3 | 1.23 | 1.28 | 1.29 | 0.85 | 1 | 0.18 | 1.33 | 1.29 | 0.28 | 1.5 | < 0.2 | 1.21 | 0.71 | 1.05 | 1.12 | 1.31
| Mean |  | 1.205 | 1.7459 | 1.4417 | 1.265 | 1.24 | 1.3 | 1.2867 | 0.88167 | 0.96667 | 0.15833 | 1.3767 | 1.2167 | 0.25333 | 1.4667 | 0.7 | 1.2083 | 1.0017 | 1.0433 | 1.17 | 1.3467
| Median |  | 1.205 | 1.7525 | 1.445 | 1.26 | 1.245 | 1.295 | 1.295 | 0.88 | 1 | 0.16 | 1.37 | 1.21 | 0.265 | 1.5 | 0.7 | 1.215 | 1.04 | 1.035 | 1.165 | 1.33
| Std Dev. |  | 0.053572 | 0.023965 | 0.041191 | 0.048888 | 0.017889 | 0.016733 | 0.02582 | 0.047924 | 0.15055 | 0.022286 | 0.051251 | 0.04761 | 0.072296 | 0.18619 | 0.14142 | 0.033116 | 0.15459 | 0.028752 | 0.04 | 0.043205
| Rel.Std.Dev. |  | 4.446% | 1.373% | 2.857% | 3.865% | 1.443% | 1.287% | 2.007% | 5.436% | 15.57% | 14.08% | 3.723% | 3.913% | 28.54% | 12.69% | 20.2% | 2.741% | 15.43% | 2.756% | 3.419% | 3.208%
| PDM3 |  | -1.383% | 42.89% | 17.99% | 3.527% | 1.481% | 6.391% | 5.3% | -27.84% | -20.89% | -87.04% | 12.67% | -0.4287% | -79.27% | 20.03% | -42.71% | -1.111% | -18.02% | -14.61% | -4.248% | 10.21%
| Z-Score (Absolute) |  | 0.03 | 2.27 | 0.98 | 0.22 | 0.12 | 0.37 | 0.32 | 1.41 | 1.05 | 4.49 | 0.7 | 0.02 | 4.08 | 1.39 | 2.18 | 0.02 | 0.65 | 0.72 | 0.18 | 0.57
| Analytical results for Hg in OREAS 905 (Certified Value < 0.05 ppm) | [31 empty] | Certified
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 3
|  |  | 0.4g | 20g | 0.25g | 0.5g | 0.5g | 0.25g | 0.15g | 0.5g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 0.01 | 0.034144 | 0.03 | < 0.01 | 0.02 | 0.17 | < 1 | 0.02 | 0.01 | < 5 | 0.02 | 2 | < 2 | < 0.02 | < 1 | 0.043 | < 1 | 0.02 | 0.03 | 0.01 | < 0.1 | 0.02 | < 0.01 | 0.014 | [6 empty] | 1
| 1 | 2 | 0.01 | 0.035331 | 0.03 | < 0.01 | 0.02 | 0.12 | < 1 | 0.03 | 0.02 | < 5 | 0.02 | 1 | < 2 | < 0.02 | < 1 | 0.9 | < 1 | 0.03 | 0.03 | 0.02 | < 0.1 | 0.02 | 0.01 | 0.014 | [6 empty] | 3
| 1 | 3 | 0.01 | 0.0273 | 0.03 | < 0.01 | 0.02 | 0.12 | < 1 | 0.01 | < 0.01 | < 5 | 0.02 | 1 | < 2 | < 0.02 | < 1 | 0.751 | < 1 | 0.03 | 0.03 | 0.04 | < 0.1 | 0.02 | 0.01 | 0.02 | [6 empty] | 16
| 1 | 4 | 0.01 | 0.018814 | 0.01 | < 0.01 | 0.01 | 0.09 | < 1 | 0.03 | 0.01 | < 5 | 0.02 | 1 | < 2 | < 0.02 | < 1 | 0.582 | < 1 | 0.02 | 0.03 | 0.05 | < 0.1 | 0.02 | < 0.01 | 0.011 | [6 empty] | < 0.05
| 1 | 5 | 0.02 | 0.028782 | 0.02 | < 0.01 | 0.02 | 0.09 | < 1 | < 0.01 | < 0.01 | < 5 | 0.01 | 1 | < 2 | < 0.02 | < 1 | 0.486 | < 1 | 0.03 | 0.02 | 0.03 | < 0.1 | 0.02 | < 0.01 | 0.014 | [6 empty] | 139
| 1 | 6 | 0.01 | 0.022606 | 0.03 | < 0.01 | 0.02 | 0.09 | < 1 | 0.02 | 0.02 | < 5 | 0.02 | 1 | < 2 | < 0.02 | < 1 | 0.414 | < 1 | 0.02 | 0.02 | 0.04 | < 0.1 | 0.02 | 0.01 | 0.019
| Mean |  | 0.011667 | 0.027829 | 0.025 |  | 0.018333 | 0.11333 |  | 0.022 | 0.015 |  | 0.018333 | 1.1667 |  |  |  | 0.52933 |  | 0.025 | 0.026667 | 0.031667 |  | 0.02 | 0.01 | 0.015333
| Median |  | 0.01 | 0.028041 | 0.03 |  | 0.02 | 0.105 |  | 0.02 | 0.015 |  | 0.02 | 1 |  |  |  | 0.534 |  | 0.025 | 0.03 | 0.035 |  | 0.02 | 0.01 | 0.014
| Std Dev. |  | 0.0040825 | 0.0064169 | 0.0083666 |  | 0.0040825 | 0.031411 |  | 0.0083666 | 0.0057735 |  | 0.0040825 | 0.40825 |  |  |  | 0.29699 |  | 0.0054772 | 0.005164 | 0.01472 |  | 0 | 0 | 0.0034448
| Rel.Std.Dev. |  | 34.99% | 23.06% | 33.47% |  | 22.27% | 27.72% |  | 38.03% | 38.49% |  | 22.27% | 34.99% |  |  |  | 56.11% |  | 21.91% | 19.36% | 46.48% |  | 0% | 0% | 22.47%
PDM3
| Z-Score (Absolute) |  | 1.54 | 0.57 | 0.2 | NA | 0.67 | 11.72 | NA | 0.2 | 1.11 | NA | 0.67 | 127.44 | NA | NA | NA | 66.02 | NA | 0.2 | 0.41 | 0.59 | NA | 0.46 | 1.76 | 1.07
| Analytical results for Ho in OREAS 905 (Certified Value 0.22 ppm) | [31 empty] | Certified
Ho | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 06: Lab; 07: Lab; 10: Lab; 11: Lab; 12: Lab | [6 empty] | 1
| No | No | 01: 03; 03: 09; 04: 12; 06: 13; 07: 17; 10: 22; 11: 26; 12: 30 | [6 empty] | ppm
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 06: AR*MS; 07: AR*MS; 10: AR*OES/MS; 11: AR*MS; 12: AR*MS | [6 empty] | 2
|  |  | 01: 20g; 03: 40g; 04: 0.5g; 06: 25g; 07: 0.5g; 10: 25g; 11: 40g; 12: 0.4g | [6 empty] | 2
| 1 | 1 | 01: 0.29655; 03: 0.22; 04: 0.2; 06: 0.18; 07: 0.26; 10: < 0.1; 11: 0.19; 12: 0.19 | [6 empty] | 1
| 1 | 2 | 01: 0.27473; 03: 0.22; 04: 0.2; 06: 0.19; 07: 0.25; 10: < 0.1; 11: 0.22; 12: 0.19 | [6 empty] | 0
| 1 | 3 | 01: 0.28224; 03: 0.22; 04: 0.2; 06: 0.19; 07: 0.26; 10: < 0.1; 11: 0.205; 12: 0.19 | [6 empty] | 16
| 1 | 4 | 01: 0.28071; 03: 0.22; 04: 0.2; 06: 0.19; 07: 0.25; 10: < 0.1; 11: 0.23; 12: 0.2 | [6 empty] | 0.22258
| 1 | 5 | 01: 0.27086; 03: 0.22; 04: 0.2; 06: 0.19; 07: 0.25; 10: < 0.1; 11: 0.19; 12: 0.19 | [6 empty] | 140
| 1 | 6 | 01: 0.29847; 03: 0.215; 04: 0.3; 06: 0.17; 07: 0.25; 10: < 0.1; 11: 0.215; 12: 0.19
| Mean |  | 01: 0.28393; 03: 0.21917; 04: 0.21667; 06: 0.185; 07: 0.25333; 11: 0.20833; 12: 0.19167
| Median |  | 01: 0.28148; 03: 0.22; 04: 0.2; 06: 0.19; 07: 0.25; 11: 0.21; 12: 0.19
| Std Dev. |  | 01: 0.011308; 03: 0.0020412; 04: 0.040825; 06: 0.0083666; 07: 0.005164; 11: 0.01633; 12: 0.0040825
| Rel.Std.Dev. |  | 01: 3.983%; 03: 0.9314%; 04: 18.84%; 06: 4.522%; 07: 2.038%; 11: 7.838%; 12: 2.13%
| PDM3 |  | 01: 27.56%; 03: -1.536%; 04: -2.659%; 06: -16.89%; 07: 13.81%; 11: -6.403%; 12: -13.89%
| Z-Score (Absolute) |  | 01: 1.81; 03: 0.07; 04: 0; 06: 0.85; 07: 0.99; 10: NA; 11: 0.22; 12: 0.67
| Analytical results for In in OREAS 905 (Certified Value 0.58 ppm) | [31 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 15 | 17 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 0.25g | 0.15g | 0.5g | 25g | 0.25g | 0.5g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [8 empty] | 3
| 1 | 1 | 0.606 | 0.46227 | 0.64 | 0.6 | 0.533 | 0.626 | 0.47 | 0.59 | < 10 | 0.57 | 0.53 | 0.59 | 0.6 | < 5 | < 0.2 | 0.62 | 0.55 | 0.615 | 0.629 | 0.24 | 0.59 | 0.62 | [8 empty] | 1
| 1 | 2 | 0.551 | 0.43274 | 0.59 | 0.6 | 0.515 | 0.632 | 0.51 | 0.61 | < 10 | 0.47 | 0.55 | 0.58 | 0.6 | < 5 | 0.2 | 0.65 | 0.555 | 0.583 | 0.532 | 0.23 | 0.554 | 0.59 | [8 empty] | 0
| 1 | 3 | 0.606 | 0.43338 | 0.61 | 0.59 | 0.528 | 0.615 | 0.49 | 0.6 | < 10 | 0.48 | 0.56 | 0.57 | 0.57 | < 5 | 0.2 | 0.66 | 0.48 | 0.588 | 0.577 | 0.24 | 0.583 | 0.6 | [8 empty] | 16
| 1 | 4 | 0.595 | 0.4185 | 0.62 | 0.58 | 0.551 | 0.628 | 0.52 | 0.6 | < 10 | 0.42 | 0.56 | 0.59 | 0.58 | < 5 | 0.3 | 0.64 | 0.62 | 0.558 | 0.633 | 0.24 | 0.591 | 0.58 | [8 empty] | 0.5788
| 1 | 5 | 0.58 | 0.42201 | 0.61 | 0.62 | 0.521 | 0.61 | 0.48 | 0.62 | < 10 | 0.56 | 0.57 | 0.54 | 0.56 | < 5 | 0.2 | 0.64 | 0.57 | 0.596 | 0.497 | 0.23 | 0.574 | 0.6 | [8 empty] | 141
| 1 | 6 | 0.598 | 0.42623 | 0.61 | 0.58 | 0.529 | 0.631 | 0.5 | 0.59 | < 10 | 0.58 | 0.57 | 0.58 | 0.61 | < 5 | 0.2 | 0.65 | 0.585 | 0.586 | 0.541 | 0.24 | 0.578 | 0.58
| Mean |  | 0.58933 | 0.43252 | 0.61333 | 0.595 | 0.5295 | 0.62367 | 0.495 | 0.60167 |  | 0.51333 | 0.55667 | 0.575 | 0.58667 |  | 0.22 | 0.64333 | 0.56 | 0.58767 | 0.56817 | 0.23667 | 0.57833 | 0.595
| Median |  | 0.5965 | 0.42948 | 0.61 | 0.595 | 0.5285 | 0.627 | 0.495 | 0.6 |  | 0.52 | 0.56 | 0.58 | 0.59 |  | 0.2 | 0.645 | 0.5625 | 0.587 | 0.559 | 0.24 | 0.5805 | 0.595
| Std Dev. |  | 0.021068 | 0.015701 | 0.01633 | 0.015166 | 0.012325 | 0.009048 | 0.018708 | 0.01169 |  | 0.065625 | 0.015055 | 0.018708 | 0.019664 |  | 0.044721 | 0.013663 | 0.046583 | 0.018554 | 0.054941 | 0.005164 | 0.013633 | 0.015166
| Rel.Std.Dev. |  | 3.575% | 3.63% | 2.662% | 2.549% | 2.328% | 1.451% | 3.779% | 1.943% |  | 12.78% | 2.705% | 3.254% | 3.352% |  | 20.33% | 2.124% | 8.318% | 3.157% | 9.67% | 2.182% | 2.357% | 2.549%
| PDM3 |  | 1.819% | -25.27% | 5.966% | 2.798% | -8.518% | 7.751% | -14.48% | 3.95% |  | -11.31% | -3.825% | -0.6572% | 1.358% |  | -61.99% | 11.15% | -3.249% | 1.531% | -1.838% | -59.11% | -0.0813% | 2.798%
| Z-Score (Absolute) |  | 0.5 | 4.6 | 0.83 | 0.44 | 1.52 | 1.3 | 2.55 | 0.64 | NA | 1.44 | 0.7 | 0.05 | 0.19 | NA | 10.79 | 1.89 | 0.6 | 0.22 | 0.36 | 10.29 | 0.05 | 0.44
| Analytical results for Ir in OREAS 905 (Indicative Value < 0.01 ppm) | [31 empty] | Indicative
Ir | Round | Replicate | 01: Lab | [6 empty] | 1
| No | No | 01: 22 | [6 empty] | ppm
|  |  | 01: AR*OES/MS | [6 empty] | 3
|  |  | 01: 25g | [6 empty] | 3
| 1 | 1 | 01: < 0.01 | [6 empty] | 1
| 1 | 2 | 01: < 0.01 | [6 empty] | 14
| 1 | 3 | 01: < 0.01 | [6 empty] | 16
| 1 | 4 | 01: < 0.01 | [6 empty] | < 0.01
| 1 | 5 | 01: < 0.01 | [6 empty] | 20
| 1 | 6 | 01: < 0.01
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for K in OREAS 905 (Certified Value 0.322 wt.%) | [31 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.34 | 0.4396 | 0.34 | 0.33 | 0.3 | 0.33 | 0.408 | 0.35 | 0.32 | 0.34 | 0.31 | 0.3057 | 0.28 | 0.33 | 0.35 | 0.32 | 0.3 | 0.352 | 0.28 | 0.43 | 0.31 | 0.3 | 0.3 | 0.29 | 0.347 | 0.31 | 0.335 | 0.31 |  |  | 1
| 1 | 2 | 0.33 | 0.44369 | 0.34 | 0.338 | 0.31 | 0.33 | 0.414 | 0.35 | 0.34 | 0.33 | 0.31 | 0.312 | 0.27 | 0.31 | 0.36 | 0.31 | 0.3 | 0.348 | 0.3 | 0.36 | 0.32 | 0.31 | 0.3 | 0.29 | 0.348 | 0.3 | 0.316 | 0.3 |  |  | 0
| 1 | 3 | 0.34 | 0.43386 | 0.34 | 0.328 | 0.31 | 0.32 | 0.418 | 0.35 | 0.33 | 0.33 | 0.31 | 0.3172 | 0.28 | 0.33 | 0.35 | 0.33 | 0.3 | 0.343 | 0.3 | 0.37 | 0.32 | 0.31 | 0.3 | 0.29 | 0.354 | 0.31 | 0.317 | 0.31 |  |  | 16
| 1 | 4 | 0.34 | 0.43397 | 0.33 | 0.331 | 0.3 | 0.32 | 0.419 | 0.36 | 0.33 | 0.35 | 0.32 | 0.3163 | 0.28 | 0.33 | 0.35 | 0.32 | 0.3 | 0.357 | 0.3 | 0.38 | 0.31 | 0.3 | 0.3 | 0.29 | 0.364 | 0.32 | 0.313 | 0.31 |  |  | 0.32201
| 1 | 5 | 0.35 | 0.44905 | 0.33 | 0.327 | 0.31 | 0.32 | 0.415 | 0.35 | 0.34 | 0.33 | 0.31 | 0.3244 | 0.28 | 0.33 | 0.36 | 0.31 | 0.3 | 0.352 | 0.3 | 0.38 | 0.31 | 0.31 | 0.29 | 0.28 | 0.349 | 0.32 | 0.32 | 0.31 |  |  | 142
| 1 | 6 | 0.35 | 0.44613 | 0.33 | 0.336 | 0.3 | 0.32 | 0.409 | 0.35 | 0.33 | 0.34 | 0.31 | 0.2886 | 0.28 | 0.32 | 0.37 | 0.32 | 0.3 | 0.352 | 0.29 | 0.38 | 0.31 | 0.31 | 0.3 | 0.29 | 0.35 | 0.31 | 0.303 | 0.32
| Mean |  | 0.34167 | 0.44105 | 0.335 | 0.33167 | 0.305 | 0.32333 | 0.41383 | 0.35167 | 0.33167 | 0.33667 | 0.31167 | 0.3107 | 0.27833 | 0.325 | 0.35667 | 0.31833 | 0.3 | 0.35067 | 0.295 | 0.38333 | 0.31333 | 0.30667 | 0.29833 | 0.28833 | 0.352 | 0.31167 | 0.31733 | 0.31
| Median |  | 0.34 | 0.44164 | 0.335 | 0.3305 | 0.305 | 0.32 | 0.4145 | 0.35 | 0.33 | 0.335 | 0.31 | 0.31415 | 0.28 | 0.33 | 0.355 | 0.32 | 0.3 | 0.352 | 0.3 | 0.38 | 0.31 | 0.31 | 0.3 | 0.29 | 0.3495 | 0.31 | 0.3165 | 0.31
| Std Dev. |  | 0.0075277 | 0.0063367 | 0.0054772 | 0.0044121 | 0.0054772 | 0.005164 | 0.004535 | 0.0040825 | 0.0075277 | 0.008165 | 0.0040825 | 0.012461 | 0.0040825 | 0.0083666 | 0.008165 | 0.0075277 | 0 | 0.0047188 | 0.0083666 | 0.024221 | 0.005164 | 0.005164 | 0.0040825 | 0.0040825 | 0.0063561 | 0.0075277 | 0.010443 | 0.0063246
| Rel.Std.Dev. |  | 2.203% | 1.437% | 1.635% | 1.33% | 1.796% | 1.597% | 1.096% | 1.161% | 2.27% | 2.425% | 1.31% | 4.011% | 1.467% | 2.574% | 2.289% | 2.365% | 0% | 1.346% | 2.836% | 6.319% | 1.648% | 1.684% | 1.368% | 1.416% | 1.806% | 2.415% | 3.291% | 2.04%
| PDM3 |  | 6.104% | 36.97% | 4.034% | 2.998% | -5.283% | 0.4105% | 28.52% | 9.209% | 2.998% | 4.551% | -3.213% | -3.513% | -13.56% | 0.9281% | 10.76% | -1.142% | -6.836% | 8.899% | -8.388% | 19.04% | -2.695% | -4.765% | -7.353% | -10.46% | 9.313% | -3.213% | -1.453% | -3.73%
| Z-Score (Absolute) |  | 0.89 | 5.12 | 0.6 | 0.46 | 0.67 | 0.11 | 3.96 | 1.31 | 0.46 | 0.67 | 0.39 | 0.43 | 1.81 | 0.18 | 1.53 | 0.11 | 0.89 | 1.27 | 1.1 | 2.26 | 0.32 | 0.6 | 0.96 | 1.38 | 1.23 | 0.39 | 0.15 | 0.46
| Analytical results for La in OREAS 905 (Certified Value 39.8 ppm) | [31 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 39.8 | 42.7773 | 40.2 | 44.5 | 37.3 | 41.8 | 37.5 | 41.2 | 41.3 | 39 | 37 | 37.09 | 38 | 38.6 | 36 | 42.23 | 36.9 | 42.3 | 36.985 | 35 | 41 | 39 | 41.3 | 41.3 | 44.2 | 31.7 | 40.7 | 41.3 | 41.4 |  | 1
| 1 | 2 | 36.1 | 41.5321 | 38.6 | 43.8 | 36.5 | 41.5 | 37.8 | 40.4 | 42.6 | 40 | 36 | 38.59 | 37 | 38.3 | 36 | 41.98 | 35.8 | 41.9 | 40.69 | 36 | 44 | 40 | 40.4 | 39.7 | 42.3 | 31.5 | 37.4 | 41 | 41.8 |  | 0
| 1 | 3 | 39.1 | 41.7075 | 39.3 | 42.8 | 37.1 | 41 | 37.7 | 41 | 40.4 | 40 | 36 | 39.51 | 37 | 39 | 36 | 43.08 | 34.1 | 41.9 | 40.755 | 37 | 44 | 40 | 36.2 | 39.2 | 44.1 | 31.8 | 39.6 | 41.8 | 40.6 |  | 16
| 1 | 4 | 38.6 | 42.278 | 38 | 42.6 | 37.9 | 41.9 | 37.9 | 41.6 | 41 | 40 | 37 | 38.18 | 37 | 38.8 | 37 | 40.85 | 35.6 | 43 | 39.13 | 37 | 44 | 38 | 42.6 | 36.6 | 47 | 32.6 | 39.9 | 41.8 | 41.6 |  | 39.8025
| 1 | 5 | 38.8 | 41.8882 | 38.8 | 42.2 | 36.7 | 40.5 | 38 | 40.8 | 42.2 | 39 | 37 | 38.55 | 37 | 39.5 | 38 | 40.84 | 35.1 | 42.8 | 39.845 | 37 | 44 | 39 | 42.8 | 38.3 | 40.2 | 32.2 | 38.6 | 41.2 | 40.3 |  | 143
| 1 | 6 | 39.2 | 42.8734 | 39.2 | 44.4 | 36.6 | 41.1 | 37.9 | 41.8 | 40.3 | 40 | 37 | 36.51 | 37 | 37.5 | 38 | 42.14 | 35.5 | 42.9 | 41.73 | 36 | 44 | 40 | 40.9 | 38.9 | 33.1 | 31.8 | 38.8 | 39.5 | 41.1
| Mean |  | 38.6 | 42.1761 | 39.0167 | 43.3833 | 37.0167 | 41.3 | 37.8 | 41.1333 | 41.3 | 39.6667 | 36.6667 | 38.0717 | 37.1667 | 38.6167 | 36.8333 | 41.8533 | 35.5 | 42.4667 | 39.8558 | 36.3333 | 43.5 | 39.3333 | 40.7 | 39 | 41.8167 | 31.9333 | 39.1667 | 41.1 | 41.1333
| Median |  | 38.95 | 42.0831 | 39 | 43.3 | 36.9 | 41.3 | 37.85 | 41.1 | 41.15 | 40 | 37 | 38.365 | 37 | 38.7 | 36.5 | 42.06 | 35.55 | 42.55 | 40.2675 | 36.5 | 44 | 39.5 | 41.1 | 39.05 | 43.2 | 31.8 | 39.2 | 41.25 | 41.25
| Std Dev. |  | 1.2915 | 0.5613 | 0.74409 | 0.98065 | 0.53072 | 0.53292 | 0.17889 | 0.5164 | 0.93808 | 0.5164 | 0.5164 | 1.0937 | 0.40825 | 0.67946 | 0.98319 | 0.86929 | 0.91433 | 0.50067 | 1.6604 | 0.8165 | 1.2247 | 0.8165 | 2.3983 | 1.5543 | 4.8288 | 0.39833 | 1.1535 | 0.84853 | 0.58538
| Rel.Std.Dev. |  | 3.346% | 1.331% | 1.907% | 2.26% | 1.434% | 1.29% | 0.4732% | 1.255% | 2.271% | 1.302% | 1.408% | 2.873% | 1.098% | 1.76% | 2.669% | 2.077% | 2.576% | 1.179% | 4.166% | 2.247% | 2.816% | 2.076% | 5.893% | 3.986% | 11.55% | 1.247% | 2.945% | 2.065% | 1.423%
| PDM3 |  | -3.021% | 5.963% | -1.974% | 8.996% | -6.999% | 3.762% | -5.031% | 3.344% | 3.762% | -0.3413% | -7.879% | -4.349% | -6.622% | -2.979% | -7.46% | 5.153% | -10.81% | 6.693% | 0.134% | -8.716% | 9.29% | -1.179% | 2.255% | -2.016% | 5.06% | -19.77% | -1.597% | 3.26% | 3.344%
| Z-Score (Absolute) |  | 0.08 | 0.92 | 0.1 | 1.31 | 0.75 | 0.64 | 0.5 | 0.58 | 0.64 | 0.11 | 0.86 | 0.41 | 0.7 | 0.23 | 0.81 | 0.81 | 1.24 | 1.01 | 0.17 | 0.97 | 1.51 | 0 | 0.73 | 0.11 | 1.37 | 2.39 | 0.05 | 0.67 | 0.58
| Analytical results for Li in OREAS 905 (Certified Value 4.87 ppm) | [31 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | [8 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.5g | 0.5g | 20g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.2g | 0.5g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [8 empty] | 2
| 1 | 1 | 5.3 | 5.8951 | 5 | 4.8 | 5.5 | 5.5 | 5 | 5 | 4.4 | 4 | 5 | 6 | 5 | 4 | 4.1 | 6.4 | 3.6 | 4.9 | 4.7 | 4 | 4.6 | 5.47 | [8 empty] | 1
| 1 | 2 | 4.9 | 5.6403 | 6 | 4.6 | 5.4 | 5.5 | 6 | 5 | 4.4 | 4.1 | 4 | 6 | 5.1 | 4 | 4.1 | 4.9 | 3.6 | 4.6 | 4.5 | 4 | 4.5 | 5.29 | [8 empty] | 0
| 1 | 3 | 5.1 | 5.8382 | 6 | 4.7 | 5.3 | 5.5 | 5 | 5 | 4.4 | 4.3 | 4 | 6 | 5.1 | 4.1 | 4 | 5.3 | 3 | 4.7 | 4.9 | 4 | 4.7 | 5.27 | [8 empty] | 16
| 1 | 4 | 6.2 | 5.5068 | 6 | 4.8 | 5.4 | 5.5 | 5 | 5 | 4.4 | 4.2 | 4 | 5 | 5.1 | 4.2 | 4.2 | 5.7 | 3.8 | 4.4 | 5 | 4 | 4.8 | 5.18 | [8 empty] | 4.8746
| 1 | 5 | 4.9 | 5.8004 | 6 | 4.6 | 5.1 | 5.6 | 5 | 5 | 4.5 | 4.6 | 5 | 6 | 5.1 | 4.2 | 4.8 | 6.2 | 3.6 | 4.5 | 4.4 | 4 | 4.7 | 5.28 | [8 empty] | 144
| 1 | 6 | 5.1 | 5.9517 | 6 | 4.5 | 5.1 | 5.2 | 5 | 5 | 4.4 | 4.2 | 5 | 6 | 5.1 | 4.2 | 5 | 5.6 | 3.7 | 4.6 | 3.5 | 4 | 4.6 | 5.08
| Mean |  | 5.25 | 5.7721 | 5.8333 | 4.6667 | 5.3 | 5.4667 | 5.1667 | 5 | 4.4167 | 4.2333 | 4.5 | 5.8333 | 5.0833 | 4.1167 | 4.3667 | 5.6833 | 3.55 | 4.6167 | 4.5 | 4 | 4.65 | 5.2617
| Median |  | 5.1 | 5.8193 | 6 | 4.65 | 5.35 | 5.5 | 5 | 5 | 4.4 | 4.2 | 4.5 | 6 | 5.1 | 4.15 | 4.15 | 5.65 | 3.6 | 4.6 | 4.6 | 4 | 4.65 | 5.275
| Std Dev. |  | 0.48888 | 0.16751 | 0.40825 | 0.12111 | 0.16733 | 0.13663 | 0.40825 | 0 | 0.040825 | 0.20656 | 0.54772 | 0.40825 | 0.040825 | 0.098319 | 0.42269 | 0.55648 | 0.28107 | 0.17224 | 0.54037 | 0 | 0.10488 | 0.12983
| Rel.Std.Dev. |  | 9.312% | 2.902% | 6.999% | 2.595% | 3.157% | 2.499% | 7.902% | 0% | 0.9243% | 4.879% | 12.17% | 6.999% | 0.8031% | 2.388% | 9.68% | 9.791% | 7.917% | 3.731% | 12.01% | 0% | 2.256% | 2.468%
| PDM3 |  | 7.702% | 18.41% | 19.67% | -4.265% | 8.728% | 12.15% | 5.992% | 2.573% | -9.394% | -13.15% | -7.684% | 19.67% | 4.283% | -15.55% | -10.42% | 16.59% | -27.17% | -5.291% | -7.684% | -17.94% | -4.607% | 7.941%
| Z-Score (Absolute) |  | 0.34 | 1.41 | 1.51 | 0.25 | 0.7 | 1.03 | 0.5 | 0.25 | 0.63 | 0.9 | 0.5 | 1.51 | 0.38 | 1.08 | 0.7 | 1.28 | 1.77 | 0.33 | 0.5 | 1.26 | 0.28 | 0.65
| Analytical results for Lu in OREAS 905 (Certified Value 0.034 ppm) | [31 empty] | Certified
Lu | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab |  | 1
| No | No | 01: 03; 03: 04; 04: 05; 05: 09; 06: 10; 07: 12; 08: 13; 09: 15; 10: 17; 11: 26; 12: 30; 13: 32 |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS; 12: AR*MS; 13: AR*OES/MS |  | 3
|  |  | 01: 20g; 03: 0.25g; 04: 0.4g; 05: 40g; 06: 0.25g; 07: 0.5g; 08: 25g; 09: 0.25g; 10: 0.5g; 11: 40g; 12: 0.4g; 13: 0.25g |  | 3
| 1 | 1 | 01: 0.029311; 03: 0.04; 04: 0.04; 05: 0.03; 06: 0.04; 07: < 0.1; 08: 0.02; 09: 0.04; 10: 0.033; 11: 0.03; 12: 0.06; 13: 0.04 |  | 1
| 1 | 2 | 01: 0.021437; 03: 0.03; 04: 0.04; 05: 0.03; 06: 0.04; 07: < 0.1; 08: 0.02; 09: 0.04; 10: 0.032; 11: 0.03; 12: 0.06; 13: 0.03 |  | 6
| 1 | 3 | 01: 0.037871; 03: 0.03; 04: 0.04; 05: 0.025; 06: 0.04; 07: < 0.1; 08: 0.02; 09: 0.04; 10: 0.038; 11: 0.025; 12: 0.06; 13: 0.04 |  | 16
| 1 | 4 | 01: 0.034071; 03: 0.03; 04: 0.04; 05: 0.03; 06: 0.03; 07: < 0.1; 08: 0.02; 09: 0.04; 10: 0.035; 11: 0.03; 12: 0.06; 13: 0.03 |  | 0.033944
| 1 | 5 | 01: 0.03292; 03: 0.03; 04: 0.04; 05: 0.03; 06: 0.04; 07: < 0.1; 08: 0.02; 09: 0.04; 10: 0.031; 11: 0.03; 12: 0.06; 13: 0.04 |  | 145
| 1 | 6 | 01: 0.026372; 03: 0.03; 04: 0.04; 05: 0.025; 06: 0.03; 07: < 0.1; 08: 0.02; 09: 0.04; 10: 0.032; 11: 0.03; 12: 0.06; 13: 0.03
| Mean |  | 01: 0.03033; 03: 0.031667; 04: 0.04; 05: 0.028333; 06: 0.036667; 08: 0.02; 09: 0.04; 10: 0.0335; 11: 0.029167; 12: 0.06; 13: 0.035
| Median |  | 01: 0.031116; 03: 0.03; 04: 0.04; 05: 0.03; 06: 0.04; 08: 0.02; 09: 0.04; 10: 0.0325; 11: 0.03; 12: 0.06; 13: 0.035
| Std Dev. |  | 01: 0.0058883; 03: 0.0040825; 04: 0; 05: 0.002582; 06: 0.005164; 08: 0; 09: 0; 10: 0.0025884; 11: 0.0020412; 12: 0; 13: 0.0054772
| Rel.Std.Dev. |  | 01: 19.41%; 03: 12.89%; 04: 0%; 05: 9.113%; 06: 14.08%; 08: 0%; 09: 0%; 10: 7.727%; 11: 6.999%; 12: 0%; 13: 15.65%
| PDM3 |  | 01: -10.65%; 03: -6.709%; 04: 17.84%; 05: -16.53%; 06: 8.021%; 08: -41.08%; 09: 17.84%; 10: -1.308%; 11: -14.07%; 12: 76.76%; 13: 3.111%
| Z-Score (Absolute) |  | 01: 0.61; 03: 0.35; 04: 1.25; 05: 1; 06: 0.61; 07: NA; 08: 2.6; 09: 1.25; 10: 0; 11: 0.67; 12: 5.11; 13: 0.29
| Analytical results for Mg in OREAS 905 (Certified Value 0.158 wt.%) | [31 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.15 | 0.16548 | 0.15 | 0.178 | 0.15 | 0.15 | 0.17 | 0.19 | 0.16 | 0.18 | 0.16 | 0.15 | 0.15 | 0.17 | 0.15 | 0.15 | 0.14 | 0.167 | 0.0876 | 0.15 | 0.1 | 0.16 | 0.16 | 0.14 | 0.14 | 0.195 | 0.15 | 0.17 | 0.15 |  | 1
| 1 | 2 | 0.14 | 0.16381 | 0.15 | 0.179 | 0.15 | 0.15 | 0.17 | 0.2 | 0.17 | 0.18 | 0.16 | 0.16 | 0.15 | 0.17 | 0.15 | 0.15 | 0.14 | 0.167 | 0.09705 | 0.16 | 0.17 | 0.17 | 0.17 | 0.14 | 0.14 | 0.194 | 0.14 | 0.162 | 0.15 |  | 0
| 1 | 3 | 0.15 | 0.1642 | 0.15 | 0.17 | 0.15 | 0.15 | 0.17 | 0.2 | 0.17 | 0.18 | 0.16 | 0.16 | 0.16 | 0.17 | 0.15 | 0.15 | 0.14 | 0.167 | 0.09995 | 0.16 | 0.17 | 0.17 | 0.17 | 0.14 | 0.14 | 0.197 | 0.15 | 0.161 | 0.15 |  | 16
| 1 | 4 | 0.15 | 0.16219 | 0.15 | 0.175 | 0.15 | 0.15 | 0.17 | 0.2 | 0.17 | 0.18 | 0.16 | 0.15 | 0.16 | 0.18 | 0.15 | 0.15 | 0.14 | 0.173 | 0.0938 | 0.16 | 0.18 | 0.16 | 0.16 | 0.14 | 0.14 | 0.198 | 0.15 | 0.16 | 0.16 |  | 0.15831
| 1 | 5 | 0.15 | 0.16611 | 0.15 | 0.173 | 0.15 | 0.15 | 0.18 | 0.2 | 0.17 | 0.18 | 0.16 | 0.16 | 0.15 | 0.16 | 0.16 | 0.15 | 0.14 | 0.178 | 0.09585 | 0.16 | 0.18 | 0.17 | 0.16 | 0.14 | 0.14 | 0.198 | 0.15 | 0.164 | 0.15 |  | 146
| 1 | 6 | 0.15 | 0.16607 | 0.15 | 0.176 | 0.15 | 0.15 | 0.17 | 0.2 | 0.17 | 0.18 | 0.16 | 0.15 | 0.15 | 0.16 | 0.16 | 0.15 | 0.14 | 0.17 | 0.111 | 0.16 | 0.18 | 0.17 | 0.17 | 0.14 | 0.14 | 0.196 | 0.15 | 0.161 | 0.15
| Mean |  | 0.14833 | 0.16464 | 0.15 | 0.17517 | 0.15 | 0.15 | 0.17167 | 0.19833 | 0.16833 | 0.18 | 0.16 | 0.155 | 0.15333 | 0.16833 | 0.15333 | 0.15 | 0.14 | 0.17033 | 0.097542 | 0.15833 | 0.16333 | 0.16667 | 0.165 | 0.14 | 0.14 | 0.19633 | 0.14833 | 0.163 | 0.15167
| Median |  | 0.15 | 0.16484 | 0.15 | 0.1755 | 0.15 | 0.15 | 0.17 | 0.2 | 0.17 | 0.18 | 0.16 | 0.155 | 0.15 | 0.17 | 0.15 | 0.15 | 0.14 | 0.1685 | 0.09645 | 0.16 | 0.175 | 0.17 | 0.165 | 0.14 | 0.14 | 0.1965 | 0.15 | 0.1615 | 0.15
| Std Dev. |  | 0.0040825 | 0.0015374 | 0 | 0.0033116 | 0 | 0 | 0.0040825 | 0.0040825 | 0.0040825 | 3.0405e-17 | 0 | 0.0054772 | 0.005164 | 0.0075277 | 0.005164 | 0 | 0 | 0.0044572 | 0.0077829 | 0.0040825 | 0.031411 | 0.005164 | 0.0054772 | 0 | 0 | 0.001633 | 0.0040825 | 0.0036878 | 0.0040825
| Rel.Std.Dev. |  | 2.752% | 0.9338% | 0% | 1.891% | 0% | 0% | 2.378% | 2.058% | 2.425% | 1.689e-14% | 0% | 3.534% | 3.368% | 4.472% | 3.368% | 0% | 0% | 2.617% | 7.979% | 2.578% | 19.23% | 3.098% | 3.32% | 0% | 0% | 0.8317% | 2.752% | 2.262% | 2.692%
| PDM3 |  | -6.302% | 4% | -5.25% | 10.65% | -5.25% | -5.25% | 8.437% | 25.28% | 6.331% | 13.7% | 1.067% | -2.091% | -3.144% | 6.331% | -3.144% | -5.25% | -11.57% | 7.594% | -38.39% | 0.01435% | 3.173% | 5.278% | 4.225% | -11.57% | -11.57% | 24.02% | -6.302% | 2.962% | -4.197%
| Z-Score (Absolute) |  | 0.67 | 0.43 | 0.56 | 1.14 | 0.56 | 0.56 | 0.9 | 2.7 | 0.67 | 1.46 | 0.11 | 0.22 | 0.34 | 0.67 | 0.34 | 0.56 | 1.24 | 0.81 | 4.1 | 0 | 1.19 | 0.56 | 0.45 | 1.24 | 1.24 | 2.56 | 0.67 | 0.22 | 0.45
| Analytical results for Mn in OREAS 905 (Certified Value 0.035 wt.%) | [31 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.0352 | 0.036388 | 0.0364 | 0.0349 | 0.0352 | 0.0342 | 0.0336 | 0.041 | 0.0354 | 0.0376 | 0.0365 | 0.0349 | 0.04 | 0.036 | 0.0338 | 0.0347 | 0.03082 | 0.0393 | 0.03128 | 0.0335 | 0.0384 | 0.0361 | 0.0355 | 0.0338 | 0.0336 | 0.0358 | 0.0352 | 0.0366 | 0.0358 |  | 1
| 1 | 2 | 0.0352 | 0.03564 | 0.036 | 0.035 | 0.0357 | 0.0355 | 0.0345 | 0.042 | 0.037 | 0.0374 | 0.0355 | 0.0357 | 0.04 | 0.0352 | 0.0351 | 0.0335 | 0.03094 | 0.0388 | 0.034212 | 0.0347 | 0.0367 | 0.0361 | 0.0356 | 0.0337 | 0.0338 | 0.0354 | 0.0331 | 0.0342 | 0.0345 |  | 0
| 1 | 3 | 0.0359 | 0.036326 | 0.0353 | 0.0355 | 0.0355 | 0.0349 | 0.0345 | 0.042 | 0.0358 | 0.0373 | 0.0358 | 0.0354 | 0.04 | 0.0348 | 0.0338 | 0.0359 | 0.03004 | 0.0384 | 0.034443 | 0.0357 | 0.0368 | 0.0368 | 0.0355 | 0.0337 | 0.0335 | 0.0347 | 0.0341 | 0.0352 | 0.0342 |  | 16
| 1 | 4 | 0.0355 | 0.035763 | 0.0363 | 0.0353 | 0.0345 | 0.035 | 0.0345 | 0.043 | 0.0378 | 0.037 | 0.0352 | 0.0346 | 0.04 | 0.0352 | 0.0329 | 0.0345 | 0.03098 | 0.0388 | 0.03312 | 0.0357 | 0.0379 | 0.0354 | 0.0343 | 0.0334 | 0.0338 | 0.0351 | 0.0344 | 0.0347 | 0.0353 |  | 0.03528
| 1 | 5 | 0.0355 | 0.036161 | 0.0345 | 0.0357 | 0.035 | 0.0335 | 0.0375 | 0.042 | 0.0412 | 0.0371 | 0.0365 | 0.0362 | 0.04 | 0.0351 | 0.0351 | 0.0344 | 0.03119 | 0.0381 | 0.033638 | 0.0357 | 0.0374 | 0.0358 | 0.0355 | 0.0333 | 0.0331 | 0.0358 | 0.0346 | 0.0354 | 0.0345 |  | 147
| 1 | 6 | 0.0362 | 0.036653 | 0.0361 | 0.0351 | 0.035 | 0.0344 | 0.0354 | 0.042 | 0.0384 | 0.0376 | 0.0355 | 0.0353 | 0.04 | 0.0342 | 0.0345 | 0.0351 | 0.03168 | 0.0391 | 0.034673 | 0.0346 | 0.0381 | 0.0353 | 0.036 | 0.0332 | 0.0337 | 0.0351 | 0.0354 | 0.0355 | 0.0367
| Mean |  | 0.035583 | 0.036155 | 0.035767 | 0.03525 | 0.03515 | 0.034583 | 0.035 | 0.042 | 0.0376 | 0.037333 | 0.035833 | 0.03535 | 0.04 | 0.035083 | 0.0342 | 0.034683 | 0.030942 | 0.03875 | 0.033561 | 0.034983 | 0.03755 | 0.035917 | 0.0354 | 0.033517 | 0.033583 | 0.035317 | 0.034467 | 0.035267 | 0.035167
| Median |  | 0.0355 | 0.036243 | 0.03605 | 0.0352 | 0.0351 | 0.03465 | 0.0345 | 0.042 | 0.0374 | 0.03735 | 0.03565 | 0.03535 | 0.04 | 0.03515 | 0.03415 | 0.0346 | 0.03096 | 0.0388 | 0.033925 | 0.0352 | 0.03765 | 0.03595 | 0.0355 | 0.03355 | 0.03365 | 0.03525 | 0.0345 | 0.0353 | 0.0349
| Std Dev. |  | 0.00039707 | 0.00038756 | 0.00073121 | 0.00030822 | 0.00042308 | 0.00070261 | 0.0013506 | 0.00063246 | 0.0021014 | 0.00025033 | 0.00055015 | 0.00056833 | 0 | 0.00058793 | 0.00086255 | 0.00079603 | 0.00053593 | 0.00044159 | 0.0012522 | 0.00089088 | 0.00070071 | 0.00054924 | 0.00057271 | 0.00024833 | 0.00026394 | 0.00043551 | 0.00082865 | 0.00081404 | 0.00095847
| Rel.Std.Dev. |  | 1.116% | 1.072% | 2.044% | 0.8744% | 1.204% | 2.032% | 3.859% | 1.506% | 5.589% | 0.6705% | 1.535% | 1.608% | 0% | 1.676% | 2.522% | 2.295% | 1.732% | 1.14% | 3.731% | 2.547% | 1.866% | 1.529% | 1.618% | 0.7409% | 0.7859% | 1.233% | 2.404% | 2.308% | 2.726%
| PDM3 |  | 0.8588% | 2.48% | 1.378% | -0.08604% | -0.3695% | -1.976% | -0.7946% | 19.05% | 6.575% | 5.819% | 1.567% | 0.1974% | 13.38% | -0.5584% | -3.062% | -1.692% | -12.3% | 9.834% | -4.874% | -0.8419% | 6.433% | 1.804% | 0.3391% | -4.999% | -4.81% | 0.1029% | -2.306% | -0.0388% | -0.3222%
| Z-Score (Absolute) |  | 0.28 | 0.8 | 0.67 | 0.01 | 0.1 | 0.61 | 0.69 | 6.03 | 1.44 | 1.85 | 0.51 | 0.07 | 4.24 | 0.16 | 0.95 | 0.52 | 3.87 | 3.12 | 1.12 | 0.25 | 2.04 | 0.58 | 0.32 | 1.57 | 1.51 | 0.04 | 0.72 | 0 | 0.09
| Analytical results for Mo in OREAS 905 (Certified Value 3.02 ppm) | [31 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 3.07 | 3.0997 | 3.2 | 3.3 | 2.99 | 2.98 | 2.8 | 3 | 3.66 | 3 | 4 | 3.6 | 5 | 3.06 | < 5 | 3.16 | 6 | < 2 | 2.9 | 3 | 3.6 | 3 | 3.7 | 3.4 | 3.12 | 3 | 3.27 | 3.04 | 2.89 |  | 1
| 1 | 2 | 2.91 | 2.7806 | 3 | 3.4 | 2.92 | 2.95 | 2.9 | 2.8 | 3.46 | 3 | 3 | 3.1 | 5 | 2.92 | < 5 | 3.08 | 3 | < 2 | 2.8 | 3 | 3.3 | 3 | 3.6 | 3.31 | 3.27 | 3 | 3.15 | 2.96 | 2.92 |  | 7
| 1 | 3 | 3.14 | 2.6979 | 3.1 | 3.2 | 2.94 | 2.96 | 2.9 | 2.8 | 2.92 | 3 | 3 | 3 | 5 | 2.97 | < 5 | 3.1 | 3 | < 2 | 2.8 | 3 | 3.3 | 3 | 3.2 | 3.22 | 3.3 | 3 | 3.23 | 2.89 | 3.06 |  | 16
| 1 | 4 | 3.06 | 3.0384 | 3.1 | 3.3 | 3.04 | 2.93 | 3 | 3 | 2.62 | 3 | 3 | 3 | 5 | 2.95 | < 5 | 3.27 | 3 | < 2 | 2.7 | 3 | 3.4 | 3 | 4 | 3.14 | 3.24 | 3 | 3.33 | 2.94 | 3.04 |  | 3.0239
| 1 | 5 | 3 | 2.8671 | 3.2 | 3.2 | 2.91 | 2.91 | 2.9 | 3 | 2.33 | 3 | 4 | 3 | 5 | 2.94 | < 5 | 3.01 | 4 | < 2 | 2.7 | 3 | 3.5 | 3 | 4 | 3.15 | 3.16 | 3 | 3.19 | 3.24 | 3.03 |  | 148
| 1 | 6 | 3.06 | 2.9549 | 3 | 3.5 | 2.97 | 2.92 | 2.9 | 3 | 2.62 | 3 | 3 | 3.2 | 5 | 2.94 | < 5 | 3.07 | 3 | < 2 | 3.5 | 2 | 3.4 | 3 | 3.7 | 3.17 | 3.04 | 3 | 3.22 | 2.98 | 3.22
| Mean |  | 3.04 | 2.9064 | 3.1 | 3.3167 | 2.9617 | 2.9417 | 2.9 | 2.9333 | 2.935 | 3 | 3.3333 | 3.15 | 5 | 2.9633 |  | 3.115 | 3.6667 |  | 2.9 | 2.8333 | 3.4167 | 3 | 3.7 | 3.2317 | 3.1883 | 3 | 3.2317 | 3.0083 | 3.0267
| Median |  | 3.06 | 2.911 | 3.1 | 3.3 | 2.955 | 2.94 | 2.9 | 3 | 2.77 | 3 | 3 | 3.05 | 5 | 2.945 |  | 3.09 | 3 |  | 2.8 | 3 | 3.4 | 3 | 3.7 | 3.195 | 3.2 | 3 | 3.225 | 2.97 | 3.035
| Std Dev. |  | 0.077717 | 0.1536 | 0.089443 | 0.1169 | 0.048751 | 0.026394 | 0.063246 | 0.10328 | 0.52268 | 0 | 0.5164 | 0.23452 | 0 | 0.050067 |  | 0.090056 | 1.2111 |  | 0.30332 | 0.40825 | 0.1169 | 0 | 0.29665 | 0.10342 | 0.099281 | 0 | 0.062743 | 0.12368 | 0.11725
| Rel.Std.Dev. |  | 2.556% | 5.285% | 2.885% | 3.525% | 1.646% | 0.8973% | 2.181% | 3.521% | 17.81% | 0% | 15.49% | 7.445% | 0% | 1.69% |  | 2.891% | 33.03% |  | 10.46% | 14.41% | 3.422% | 0% | 8.018% | 3.2% | 3.114% | 0% | 1.942% | 4.111% | 3.874%
| PDM3 |  | 0.5317% | -3.885% | 2.516% | 9.681% | -2.059% | -2.72% | -4.098% | -2.996% | -2.941% | -0.7911% | 10.23% | 4.169% | 65.35% | -2.004% |  | 3.012% | 21.26% |  | -4.098% | -6.303% | 12.99% | -0.7911% | 22.36% | 6.87% | 5.437% | -0.7911% | 6.87% | -0.5155% | 0.0908%
| Z-Score (Absolute) |  | 0.41 | 0.97 | 1.04 | 3.29 | 0.4 | 0.61 | 1.04 | 0.69 | 0.67 | 0 | 0 | 0.62 | 20.75 | 0.58 | NA | 1.19 | 0 | NA | 2.28 | 0 | 4.32 | 0 | 7.26 | 2.4 | 1.95 | 0 | 2.4 | 0.39 | 0.28
| Analytical results for Na in OREAS 905 (Certified Value 0.092 wt.%) | [31 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.09 | 0.08779 | 0.09 | 0.085 | 0.09 | 0.09 | 0.11 | 0.1 | 0.09 | 0.1 | 0.091 | 0.09 | 0.08 | 0.11 | 0.11 | 0.09 | 0.1 | 0.075 | 0.09 | 0.17 | 0.08 | 0.1 | 0.09 | 0.09 | 0.077 | 0.09 | 0.09 | 0.087 |  |  | 1
| 1 | 2 | 0.09 | 0.10882 | 0.09 | 0.085 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1 | 0.1 | 0.09 | 0.09 | 0.08 | 0.1 | 0.11 | 0.09 | 0.09 | 0.075 | 0.09 | 0.08 | 0.08 | 0.1 | 0.09 | 0.09 | 0.076 | 0.09 | 0.083 | 0.087 |  |  | 8
| 1 | 3 | 0.09 | 0.1065 | 0.09 | 0.084 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1 | 0.1 | 0.089 | 0.09 | 0.08 | 0.11 | 0.1 | 0.09 | 0.09 | 0.078 | 0.1 | 0.09 | 0.08 | 0.1 | 0.09 | 0.09 | 0.077 | 0.09 | 0.081 | 0.088 |  |  | 16
| 1 | 4 | 0.09 | 0.092618 | 0.09 | 0.084 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1 | 0.11 | 0.089 | 0.09 | 0.08 | 0.11 | 0.1 | 0.09 | 0.09 | 0.078 | 0.1 | 0.11 | 0.08 | 0.1 | 0.09 | 0.09 | 0.079 | 0.09 | 0.082 | 0.09 |  |  | 0.092062
| 1 | 5 | 0.09 | 0.098402 | 0.08 | 0.0835 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1 | 0.11 | 0.09 | 0.09 | 0.08 | 0.11 | 0.1 | 0.09 | 0.1 | 0.05 | 0.1 | 0.12 | 0.08 | 0.1 | 0.09 | 0.09 | 0.079 | 0.09 | 0.081 | 0.087 |  |  | 149
| 1 | 6 | 0.09 | 0.10409 | 0.09 | 0.0855 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1 | 0.1 | 0.089 | 0.09 | 0.08 | 0.1 | 0.11 | 0.09 | 0.09 | 0.078 | 0.09 | 0.1 | 0.08 | 0.1 | 0.09 | 0.09 | 0.077 | 0.09 | 0.08 | 0.09
| Mean |  | 0.09 | 0.099704 | 0.088333 | 0.0845 | 0.09 | 0.09 | 0.11 | 0.1 | 0.098333 | 0.10333 | 0.089667 | 0.09 | 0.08 | 0.10667 | 0.105 | 0.09 | 0.093333 | 0.072333 | 0.095 | 0.11167 | 0.08 | 0.1 | 0.09 | 0.09 | 0.0775 | 0.09 | 0.082833 | 0.088167
| Median |  | 0.09 | 0.10125 | 0.09 | 0.0845 | 0.09 | 0.09 | 0.11 | 0.1 | 0.1 | 0.1 | 0.0895 | 0.09 | 0.08 | 0.11 | 0.105 | 0.09 | 0.09 | 0.0765 | 0.095 | 0.105 | 0.08 | 0.1 | 0.09 | 0.09 | 0.077 | 0.09 | 0.0815 | 0.0875
| Std Dev. |  | 1.5202e-17 | 0.0082754 | 0.0040825 | 0.0007746 | 1.5202e-17 | 1.5202e-17 | 0 | 1.5202e-17 | 0.0040825 | 0.005164 | 0.0008165 | 1.5202e-17 | 0 | 0.005164 | 0.0054772 | 1.5202e-17 | 0.005164 | 0.011039 | 0.0054772 | 0.031885 | 0 | 1.5202e-17 | 1.5202e-17 | 1.5202e-17 | 0.0012247 | 1.5202e-17 | 0.003656 | 0.001472
| Rel.Std.Dev. |  | 1.689e-14% | 8.3% | 4.622% | 0.9167% | 1.689e-14% | 1.689e-14% | 0% | 1.52e-14% | 4.152% | 4.997% | 0.9106% | 1.689e-14% | 0% | 4.841% | 5.216% | 1.689e-14% | 5.533% | 15.26% | 5.766% | 28.55% | 0% | 1.52e-14% | 1.689e-14% | 1.689e-14% | 1.58% | 1.689e-14% | 4.414% | 1.67%
| PDM3 |  | -2.24% | 8.301% | -4.05% | -8.214% | -2.24% | -2.24% | 19.48% | 8.622% | 6.812% | 12.24% | -2.602% | -2.24% | -13.1% | 15.86% | 14.05% | -2.24% | 1.381% | -21.43% | 3.191% | 21.3% | -13.1% | 8.622% | -2.24% | -2.24% | -15.82% | -2.24% | -10.02% | -4.231%
| Z-Score (Absolute) |  | 0 | 0.95 | 0.16 | 0.54 | 0 | 0 | 1.95 | 0.97 | 0.81 | 1.3 | 0.03 | 0 | 0.97 | 1.62 | 1.46 | 0 | 0.32 | 1.29 | 0.49 | 0.97 | 0.97 | 0.97 | 0 | 0 | 1.22 | 0 | 0.84 | 0.18
| Analytical results for Nb in OREAS 905 (Indicative Value 0.44 ppm) | [31 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [7 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 0.25g | 0.15g | 0.5g | 25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [7 empty] | 2
| 1 | 1 | 0.24 | 0.56937 | 0.8 | 0.5 | 0.29 | 0.37 | 0.45 | 0.58 | < 10 | 0.3 | < 0.2 | 0.7 | 0.28 | < 1 | < 5 | < 0.1 | 0.4 | < 0.2 | 0.4 | 0.29 | 1 | 0.34 | 0.62 | [7 empty] | 1
| 1 | 2 | 0.24 | 0.55428 | 0.79 | 0.5 | 0.28 | 0.37 | 0.43 | 0.86 | < 10 | 0.3 | < 0.2 | 0.73 | 0.32 | < 1 | < 5 | < 0.1 | 0.4 | < 0.2 | 0.4 | 0.29 | 1 | 0.32 | 0.58 | [7 empty] | 15
| 1 | 3 | 0.23 | 0.52195 | 0.76 | 0.6 | 0.32 | 0.36 | 0.46 | 0.47 | < 10 | 0.3 | < 0.2 | 0.74 | 0.27 | < 1 | < 5 | < 0.1 | 0.4 | < 0.2 | 0.41 | 0.32 | 1.1 | 0.31 | 0.59 | [7 empty] | 16
| 1 | 4 | 0.23 | 0.54101 | 0.77 | 0.6 | 0.3 | 0.33 | 0.44 | 0.45 | < 10 | 0.4 | < 0.2 | 0.76 | 0.27 | < 1 | < 5 | < 0.1 | 0.4 | < 0.2 | 0.38 | 0.32 | 1.1 | 0.34 | 0.6 | [7 empty] | 0.43926
| 1 | 5 | 0.24 | 0.50623 | 0.74 | 0.6 | 0.29 | 0.36 | 0.5 | 0.47 | < 10 | 0.3 | < 0.2 | 0.68 | 0.26 | < 1 | < 5 | < 0.1 | 0.4 | < 0.2 | 0.37 | 0.32 | 0.8 | 0.35 | 0.61 | [7 empty] | 21
| 1 | 6 | 0.36 | 0.53051 | 0.8 | 0.5 | 0.29 | 0.35 | 0.55 | 0.4 | < 10 | 0.4 | < 0.2 | 0.65 | 0.3 | < 1 | < 5 | < 0.1 | 0.4 | < 0.2 | 0.4 | 0.28 | 0.9 | 0.31 | 0.6
| Mean |  | 0.25667 | 0.53723 | 0.77667 | 0.55 | 0.295 | 0.35667 | 0.47167 | 0.53833 |  | 0.33333 |  | 0.71 | 0.28333 |  |  |  | 0.4 |  | 0.39333 | 0.30333 | 0.98333 | 0.32833 | 0.6
| Median |  | 0.24 | 0.53576 | 0.78 | 0.55 | 0.29 | 0.36 | 0.455 | 0.47 |  | 0.3 |  | 0.715 | 0.275 |  |  |  | 0.4 |  | 0.4 | 0.305 | 1 | 0.33 | 0.6
| Std Dev. |  | 0.050859 | 0.022705 | 0.024221 | 0.054772 | 0.013784 | 0.015055 | 0.04535 | 0.16822 |  | 0.05164 |  | 0.040988 | 0.022509 |  |  |  | 6.0809e-17 |  | 0.015055 | 0.018619 | 0.1169 | 0.017224 | 0.014142
| Rel.Std.Dev. |  | 19.82% | 4.226% | 3.119% | 9.959% | 4.673% | 4.221% | 9.615% | 31.25% |  | 15.49% |  | 5.773% | 7.944% |  |  |  | 1.52e-14% |  | 3.828% | 6.138% | 11.89% | 5.246% | 2.357%
| PDM3 |  | -41.57% | 22.3% | 76.81% | 25.21% | -32.84% | -18.8% | 7.377% | 22.55% |  | -24.12% |  | 61.63% | -35.5% |  |  |  | -8.939% |  | -10.46% | -30.95% | 123.9% | -25.25% | 36.59%
| Z-Score (Absolute) |  | 1.01 | 0.84 | 2.31 | 0.92 | 0.67 | 0.27 | 0.34 | 0.45 | NA | 0.41 | NA | 1.9 | 0.72 | NA | NA | NA | 0 | NA | 0.04 | 0.59 | 3.58 | 0.44 | 1.23
| Analytical results for Nd in OREAS 905 (Certified Value 30.5 ppm) | [31 empty] | Certified
Nd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 12; 06: 13; 07: 17; 08: 22; 09: 26; 10: 30 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*OES/MS; 09: AR*MS; 10: AR*MS |  |  | 1
|  |  | 01: 20g; 03: 0.4g; 04: 40g; 05: 0.5g; 06: 25g; 07: 0.5g; 08: 25g; 09: 40g; 10: 0.4g |  |  | 1
| 1 | 1 | 01: 35.4564; 03: 33.5; 04: 31.9; 05: 29.8; 06: 30.07; 07: 32.46; 08: 22.5; 09: 29; 10: 26.8 |  |  | 1
| 1 | 2 | 01: 35.0141; 03: 34.2; 04: 33.9; 05: 28.4; 06: 30.99; 07: 31.46; 08: 24.5; 09: 29.8; 10: 27.1 |  |  | 0
| 1 | 3 | 01: 35.3549; 03: 33.8; 04: 33; 05: 28.6; 06: 31.69; 07: 32.93; 08: 24.5; 09: 26.2; 10: 26.7 |  |  | 16
| 1 | 4 | 01: 35.9448; 03: 34; 04: 33; 05: 28.1; 06: 30.66; 07: 31.43; 08: 23.6; 09: 31.5; 10: 26.5 |  |  | 30.4811
| 1 | 5 | 01: 35.8443; 03: 35; 04: 32.4; 05: 29.1; 06: 30.91; 07: 31.32; 08: 24.1; 09: 29.6; 10: 26.3 |  |  | 150
| 1 | 6 | 01: 35.9849; 03: 33.2; 04: 31.4; 05: 29; 06: 29.54; 07: 31.98; 08: 25.2; 09: 30.3; 10: 26.6
| Mean |  | 01: 35.5999; 03: 33.95; 04: 32.6; 05: 28.8333; 06: 30.6433; 07: 31.93; 08: 24.0667; 09: 29.4; 10: 26.6667
| Median |  | 01: 35.6503; 03: 33.9; 04: 32.7; 05: 28.8; 06: 30.785; 07: 31.72; 08: 24.3; 09: 29.7; 10: 26.65
| Std Dev. |  | 01: 0.38747; 03: 0.6253; 04: 0.89219; 05: 0.60222; 06: 0.75275; 07: 0.65213; 08: 0.93095; 09: 1.7788; 10: 0.27325
| Rel.Std.Dev. |  | 01: 1.088%; 03: 1.842%; 04: 2.737%; 05: 2.089%; 06: 2.456%; 07: 2.042%; 08: 3.868%; 09: 6.05%; 10: 1.025%
| PDM3 |  | 01: 16.79%; 03: 11.38%; 04: 6.952%; 05: -5.406%; 06: 0.5322%; 07: 4.753%; 08: -21.04%; 09: -3.547%; 10: -12.51%
| Z-Score (Absolute) |  | 01: 1.71; 03: 1.14; 04: 0.67; 05: 0.62; 06: 0; 07: 0.44; 08: 2.27; 09: 0.21; 10: 1.37
| Analytical results for Ni in OREAS 905 (Certified Value 8.9 ppm) | [31 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 9.2 | 9.4614 | 9.4 | 9 | 8.4 | 9.6 | 8.1 | 10 | 9 | 9 | 10.4 | 8.8 | 8 | 10.1 | 19 | 8 | 8.3 | 11.9 | 7.3 | 9 | 7.5 | 9 | 14 | 10 | 8.2 | 9 | 8.2 | 8.61 | 9.2 |  | 1
| 1 | 2 | 8.5 | 9.6631 | 8.7 | 9 | 8.6 | 9.7 | 8 | 9 | 9 | 9 | 8.6 | 8.9 | 7 | 8.9 | 23 | 9 | 7.6 | 12.3 | 6.4 | 9 | 9.6 | 9 | 13 | 10.5 | 8.6 | 9 | 7.8 | 8.39 | 9.6 |  | 9
| 1 | 3 | 9 | 9.0627 | 9 | 9 | 8.4 | 9.9 | 8.3 | 10 | 8.3 | 11 | 8.7 | 9.3 | 8 | 9.2 | 14 | 10 | 7 | 12 | 6.5 | 9 | 12.6 | 9 | 12 | 9.5 | 9.4 | 9 | 8.2 | 8.12 | 10 |  | 16
| 1 | 4 | 9.6 | 9.1898 | 9.6 | 9 | 9.5 | 9.7 | 7.9 | 9 | 8.8 | 9 | 10.2 | 8.9 | 6 | 9.6 | 16 | 8 | 7.2 | 11.8 | 6.2 | 9 | 9.3 | 9 | 13 | 8.8 | 9.2 | 9 | 7.8 | 8.02 | 9.2 |  | 8.903
| 1 | 5 | 9 | 9.3113 | 9 | 9 | 8.5 | 8.9 | 7.8 | 9 | 9.5 | 9 | 9.4 | 9 | 7 | 8.8 | 16 | 8 | 8.5 | 11.7 | 6.2 | 9 | 9.2 | 9 | 12 | 9 | 8.4 | 9 | 8.4 | 9.48 | 9.1 |  | 151
| 1 | 6 | 9.2 | 9.4979 | 9.6 | 9 | 8.9 | 8.9 | 7.8 | 9 | 8.6 | 9 | 9.4 | 9.2 | 7 | 8.9 | 18 | 8 | 8.1 | 11.2 | 7 | 11 | 9.1 | 9 | 12 | 9.4 | 7.7 | 9 | 7.9 | 8.06 | 9.7
| Mean |  | 9.0833 | 9.3643 | 9.2167 | 9 | 8.7167 | 9.45 | 7.9833 | 9.3333 | 8.8667 | 9.3333 | 9.45 | 9.0167 | 7.1667 | 9.25 | 17.6667 | 8.5 | 7.7833 | 11.8167 | 6.6 | 9.3333 | 9.55 | 9 | 12.6667 | 9.5333 | 8.5833 | 9 | 8.05 | 8.4467 | 9.4667
| Median |  | 9.1 | 9.3863 | 9.2 | 9 | 8.55 | 9.65 | 7.95 | 9 | 8.9 | 9 | 9.4 | 8.95 | 7 | 9.05 | 17 | 8 | 7.85 | 11.85 | 6.45 | 9 | 9.25 | 9 | 12.5 | 9.45 | 8.5 | 9 | 8.05 | 8.255 | 9.4
| Std Dev. |  | 0.36009 | 0.21945 | 0.37103 | 0 | 0.42622 | 0.43704 | 0.19408 | 0.5164 | 0.40825 | 0.8165 | 0.74229 | 0.19408 | 0.75277 | 0.50892 | 3.1411 | 0.83666 | 0.61128 | 0.3656 | 0.45166 | 0.8165 | 1.667 | 0 | 0.8165 | 0.6314 | 0.63377 | 0 | 0.251 | 0.55424 | 0.3559
| Rel.Std.Dev. |  | 3.964% | 2.343% | 4.026% | 0% | 4.89% | 4.625% | 2.431% | 5.533% | 4.604% | 8.748% | 7.855% | 2.152% | 10.5% | 5.502% | 17.78% | 9.843% | 7.854% | 3.094% | 6.843% | 8.748% | 17.46% | 0% | 6.446% | 6.623% | 7.384% | 0% | 3.118% | 6.562% | 3.76%
| PDM3 |  | 2.026% | 5.182% | 3.524% | 1.09% | -2.092% | 6.144% | -10.33% | 4.834% | -0.4076% | 4.834% | 6.144% | 1.277% | -19.5% | 3.898% | 98.44% | -4.526% | -12.58% | 32.73% | -25.87% | 4.834% | 7.268% | 1.09% | 42.27% | 7.081% | -3.59% | 1.09% | -9.581% | -5.125% | 6.332%
| Z-Score (Absolute) |  | 0.13 | 0.56 | 0.33 | 0 | 0.67 | 0.69 | 1.56 | 0.51 | 0.2 | 0 | 0.69 | 0.03 | 2.81 | 0.38 | 13.28 | 0.77 | 1.86 | 4.32 | 3.68 | 0 | 0.09 | 0 | 5.62 | 0.82 | 0.64 | 0 | 1.46 | 1.16 | 0.72
| Analytical results for Os in OREAS 905 (Indicative Value < 0.01 ppm) | [31 empty] | Indicative
Os | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 22 |  | ppm
|  |  | 01: AR*OES/MS |  | 3
|  |  | 01: 25g |  | 3
| 1 | 1 | 01: < 0.01 |  | 1
| 1 | 2 | 01: 0.06 |  | 12
| 1 | 3 | 01: < 0.01 |  | 16
| 1 | 4 | 01: < 0.01 |  | < 0.01
| 1 | 5 | 01: < 0.01 |  | 18
| 1 | 6 | 01: < 0.01
| Mean |  | 01: 0.06
| Median |  | 01: 0.06
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01
| Analytical results for P in OREAS 905 (Certified Value 0.023 wt.%) | [31 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 3
| 1 | 1 | 0.024 | 0.021002 | 0.0238 | 0.0235 | 0.025 | 0.025 | 0.025 | 0.03 | 0.022 | 0.0269 | 0.022 | 0.0206 | 0.02 | 0.027 | 0.03 | 0.0236 | 0.03 | 0.025 | 0.0211 | 0.023 | 0.026 | 0.023 | 0.024 | 0.024 | 0.024 | 0.014 | 0.024 | 0.0229 | 0.024 |  | 1
| 1 | 2 | 0.024 | 0.02297 | 0.0223 | 0.0245 | 0.025 | 0.025 | 0.025 | 0.03 | 0.024 | 0.0267 | 0.021 | 0.0209 | 0.02 | 0.028 | 0.03 | 0.0227 | 0.02 | 0.0249 | 0.0228 | 0.023 | 0.029 | 0.024 | 0.024 | 0.024 | 0.024 | 0.012 | 0.023 | 0.0221 | 0.023 |  | 0
| 1 | 3 | 0.024 | 0.020145 | 0.0229 | 0.024 | 0.025 | 0.025 | 0.025 | 0.03 | 0.023 | 0.0265 | 0.021 | 0.0211 | 0.02 | 0.027 | 0.03 | 0.0238 | 0.03 | 0.0255 | 0.0231 | 0.023 | 0.03 | 0.024 | 0.022 | 0.024 | 0.024 | 0.012 | 0.024 | 0.0215 | 0.024 |  | 16
| 1 | 4 | 0.024 | 0.023602 | 0.0222 | 0.024 | 0.024 | 0.025 | 0.024 | 0.03 | 0.024 | 0.0267 | 0.022 | 0.0205 | 0.02 | 0.026 | 0.03 | 0.0231 | 0.02 | 0.0252 | 0.0222 | 0.023 | 0.03 | 0.023 | 0.024 | 0.023 | 0.024 | 0.013 | 0.024 | 0.0213 | 0.026 |  | 0.023418
| 1 | 5 | 0.024 | 0.023458 | 0.0231 | 0.0235 | 0.025 | 0.025 | 0.025 | 0.03 | 0.022 | 0.0265 | 0.022 | 0.0211 | 0.02 | 0.025 | 0.03 | 0.0229 | 0.03 | 0.0257 | 0.0226 | 0.023 | 0.03 | 0.023 | 0.024 | 0.023 | 0.023 | 0.013 | 0.024 | 0.0225 | 0.025 |  | 152
| 1 | 6 | 0.024 | 0.024347 | 0.0223 | 0.0245 | 0.024 | 0.025 | 0.024 | 0.028 | 0.024 | 0.0267 | 0.022 | 0.0215 | 0.02 | 0.027 | 0.03 | 0.0231 | 0.03 | 0.025 | 0.0237 | 0.023 | 0.029 | 0.023 | 0.024 | 0.024 | 0.024 | 0.011 | 0.024 | 0.0216 | 0.024
| Mean |  | 0.024 | 0.022587 | 0.022767 | 0.024 | 0.024667 | 0.025 | 0.024667 | 0.029667 | 0.023167 | 0.026667 | 0.021667 | 0.02095 | 0.02 | 0.026667 | 0.03 | 0.0232 | 0.026667 | 0.025217 | 0.022583 | 0.023 | 0.029 | 0.023333 | 0.023667 | 0.023667 | 0.023833 | 0.0125 | 0.023833 | 0.021983 | 0.024333
| Median |  | 0.024 | 0.023214 | 0.0226 | 0.024 | 0.025 | 0.025 | 0.025 | 0.03 | 0.0235 | 0.0267 | 0.022 | 0.021 | 0.02 | 0.027 | 0.03 | 0.0231 | 0.03 | 0.0251 | 0.0227 | 0.023 | 0.0295 | 0.023 | 0.024 | 0.024 | 0.024 | 0.0125 | 0.024 | 0.02185 | 0.024
| Std Dev. |  | 3.8006e-18 | 0.0016436 | 0.00062503 | 0.00044721 | 0.0005164 | 3.8006e-18 | 0.0005164 | 0.0008165 | 0.00098319 | 0.00015055 | 0.0005164 | 0.00036742 | 0 | 0.0010328 | 0 | 0.00041952 | 0.005164 | 0.00031885 | 0.00088412 | 3.8006e-18 | 0.0015492 | 0.0005164 | 0.0008165 | 0.0005164 | 0.00040825 | 0.0010488 | 0.00040825 | 0.00062743 | 0.0010328
| Rel.Std.Dev. |  | 1.584e-14% | 7.277% | 2.745% | 1.863% | 2.094% | 1.52e-14% | 2.094% | 2.752% | 4.244% | 0.5646% | 2.383% | 1.754% | 0% | 3.873% | 0% | 1.808% | 19.36% | 1.264% | 3.915% | 1.652e-14% | 5.342% | 2.213% | 3.45% | 2.182% | 1.713% | 8.39% | 1.713% | 2.854% | 4.244%
| PDM3 |  | 2.485% | -3.549% | -2.782% | 2.485% | 5.331% | 6.755% | 5.331% | 26.68% | -1.074% | 13.87% | -7.479% | -10.54% | -14.6% | 13.87% | 28.11% | -0.9315% | 13.87% | 7.68% | -3.565% | -1.786% | 23.84% | -0.3622% | 1.061% | 1.061% | 1.773% | -46.62% | 1.773% | -6.127% | 3.908%
| Z-Score (Absolute) |  | 0.21 | 0.67 | 0.56 | 0.21 | 0.62 | 0.83 | 0.62 | 3.75 | 0.31 | 1.87 | 1.25 | 1.7 | 2.29 | 1.87 | 3.96 | 0.29 | 2.29 | 0.97 | 0.68 | 0.42 | 3.71 | 0.21 | 0 | 0 | 0.1 | 6.97 | 0.1 | 1.05 | 0.42
| Analytical results for Pb in OREAS 905 (Certified Value 17.1 ppm) | [31 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.4g | 20g | 0.25g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 15.7 | 16.6827 | 17.3 | 16.8 | 17.5 | 15.7 | 15 | 18.6 | 20 | 16.6 | 15.7 | 16 | 16.6 | 21 | 17 | 19 | 19 | 17 | 15 | 19.9 | 20 | 18 | 18.4 | 19.2 | 18 | 14.9 | 17.2 | 18.37 |  |  | 1
| 1 | 2 | 15.5 | 16.6661 | 17.1 | 16.2 | 17.4 | 15.6 | 15 | 17.8 | 21 | 15.5 | 15.7 | 12 | 16.5 | 19 | 16.4 | 17 | 18.6 | 17 | 15 | 18.2 | 20 | 16 | 18 | 18 | 17 | 14.5 | 16.5 | 18.22 |  |  | 0
| 1 | 3 | 16.3 | 16.3924 | 16.9 | 17.2 | 17.7 | 15.5 | 16 | 16.6 | 19 | 15.9 | 15.8 | 16 | 16.7 | 23 | 17 | 19 | 18.6 | 20 | 15 | 18.5 | 20 | 19 | 18.1 | 18.8 | 18 | 15.4 | 16.8 | 15.82 |  |  | 16
| 1 | 4 | 16.2 | 16.638 | 16.6 | 17.1 | 17.9 | 15.3 | 15 | 17 | 20 | 16.3 | 14.8 | 12 | 16.9 | 18 | 16.5 | 18 | 19.8 | 14 | 16 | 18.7 | 19 | 15 | 17.9 | 19.4 | 19 | 15.1 | 17.1 | 17.51 |  |  | 17.0849
| 1 | 5 | 16.2 | 16.4116 | 16.8 | 16.2 | 17.3 | 15.1 | 16 | 17.2 | 20 | 16.3 | 15.7 | 15 | 15.7 | 22 | 16.6 | 18 | 19.1 | 15 | 17 | 18.3 | 19 | 18 | 17.8 | 16.1 | 17 | 15 | 16.8 | 16.45 |  |  | 153
| 1 | 6 | 15.8 | 16.4298 | 17.1 | 16.4 | 17.8 | 15 | 15 | 16.7 | 20 | 16.1 | 18.7 | 15 | 16.9 | 20 | 16.8 | 18 | 18.3 | 21 | 15 | 18.6 | 19 | 15 | 17.8 | 15.4 | 17 | 15.2 | 16.9 | 16.55
| Mean |  | 15.95 | 16.5368 | 16.9667 | 16.65 | 17.6 | 15.3667 | 15.3333 | 17.3167 | 20 | 16.1167 | 16.0667 | 14.3333 | 16.55 | 20.5 | 16.7167 | 18.1667 | 18.9 | 17.3333 | 15.5 | 18.7 | 19.5 | 16.8333 | 18 | 17.8167 | 17.6667 | 15.0167 | 16.8833 | 17.1533
| Median |  | 16 | 16.5339 | 17 | 16.6 | 17.6 | 15.4 | 15 | 17.1 | 20 | 16.2 | 15.7 | 15 | 16.65 | 20.5 | 16.7 | 18 | 18.8 | 17 | 15 | 18.55 | 19.5 | 17 | 17.95 | 18.4 | 17.5 | 15.05 | 16.85 | 17.03
| Std Dev. |  | 0.32711 | 0.1387 | 0.25033 | 0.44609 | 0.23664 | 0.28048 | 0.5164 | 0.76004 | 0.63246 | 0.38166 | 1.3426 | 1.8619 | 0.44609 | 1.8708 | 0.25626 | 0.75277 | 0.52915 | 2.7325 | 0.83666 | 0.61644 | 0.54772 | 1.7224 | 0.22804 | 1.6857 | 0.8165 | 0.30605 | 0.24833 | 1.0375
| Rel.Std.Dev. |  | 2.051% | 0.8388% | 1.475% | 2.679% | 1.345% | 1.825% | 3.368% | 4.389% | 3.162% | 2.368% | 8.357% | 12.99% | 2.695% | 9.126% | 1.533% | 4.144% | 2.8% | 15.76% | 5.398% | 3.296% | 2.809% | 10.23% | 1.267% | 9.462% | 4.622% | 2.038% | 1.471% | 6.048%
| PDM3 |  | -6.643% | -3.208% | -0.6919% | -2.545% | 3.015% | -10.06% | -10.25% | 1.357% | 17.06% | -5.667% | -5.96% | -16.11% | -3.131% | 19.99% | -2.155% | 6.332% | 10.62% | 1.454% | -9.277% | 9.453% | 14.14% | -1.472% | 5.356% | 4.283% | 3.405% | -12.11% | -1.18% | 0.4006%
| Z-Score (Absolute) |  | 0.7 | 0.28 | 0.03 | 0.2 | 0.49 | 1.13 | 1.15 | 0.28 | 2.22 | 0.58 | 1 | 1.87 | 0.15 | 2.22 | 0.15 | 0.9 | 1.43 | 0.3 | 1.03 | 1.11 | 1.86 | 0.07 | 0.78 | 0.64 | 0.54 | 1.38 | 0.03 | 0.16
| Analytical results for Pd in OREAS 905 (Indicative Value < 10 ppb) | [31 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 06: Lab |  |  | 1
| No | No | 01: 03; 03: 13; 04: 17; 06: 22 |  |  | ppb
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 06: AR*OES/MS |  |  | 0
|  |  | 01: 20g; 03: 25g; 04: 0.5g; 06: 25g |  |  | 0
| 1 | 1 | 01: < 10; 03: 13; 04: < 5; 06: < 10 |  |  | 1
| 1 | 2 | 01: < 10; 03: 11; 04: < 5; 06: < 10 |  |  | 1
| 1 | 3 | 01: < 10; 03: < 10; 04: < 5; 06: < 10 |  |  | 16
| 1 | 4 | 01: < 10; 03: < 10; 04: < 5; 06: < 10 |  |  | < 10
| 1 | 5 | 01: < 10; 03: < 10; 04: < 5; 06: < 10 |  |  | 19
| 1 | 6 | 01: < 10; 03: < 10; 04: < 5; 06: < 10
| Mean |  | 03: 12
| Median |  | 03: 12
| Std Dev. |  | 03: 1.4142
| Rel.Std.Dev. |  | 03: 11.79%
PDM3
Z-Score (Absolute) | bar: 01 to 06
| Analytical results for Pr in OREAS 905 (Certified Value 8.42 ppm) | [31 empty] | Certified
Pr | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab |  |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 06: 12; 07: 13; 08: 17; 09: 22; 10: 26; 11: 30 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*MS |  |  | 2
|  |  | 01: 20g; 03: 0.4g; 04: 40g; 06: 0.5g; 07: 25g; 08: 0.5g; 09: 25g; 10: 40g; 11: 0.4g |  |  | 3
| 1 | 1 | 01: 9.6564; 03: 8.44; 04: 8.43; 06: 8.2; 07: 8.39; 08: 8.64; 09: 4.6; 10: 7.75; 11: 7.28 |  |  | 1
| 1 | 2 | 01: 9.3619; 03: 8.62; 04: 8.34; 06: 8.1; 07: 8.72; 08: 8.44; 09: 5; 10: 8.47; 11: 7.27 |  |  | 0
| 1 | 3 | 01: 9.3552; 03: 8.48; 04: 8.52; 06: 8.1; 07: 8.95; 08: 8.8; 09: 5; 10: 6.57; 11: 7.26 |  |  | 16
| 1 | 4 | 01: 9.604; 03: 8.65; 04: 8.46; 06: 8.1; 07: 8.57; 08: 8.45; 09: 4.8; 10: 9.86; 11: 7.27 |  |  | 8.4199
| 1 | 5 | 01: 9.5272; 03: 8.75; 04: 8.61; 06: 8.3; 07: 8.62; 08: 8.43; 09: 4.9; 10: 7.96; 11: 7.04 |  |  | 154
| 1 | 6 | 01: 9.6165; 03: 8.29; 04: 8.33; 06: 8; 07: 8.29; 08: 8.55; 09: 5.2; 10: 9.28; 11: 7.23
| Mean |  | 01: 9.5202; 03: 8.5383; 04: 8.4483; 06: 8.1333; 07: 8.59; 08: 8.5517; 09: 4.9167; 10: 8.315; 11: 7.225
| Median |  | 01: 9.5656; 03: 8.55; 04: 8.445; 06: 8.1; 07: 8.595; 08: 8.5; 09: 4.95; 10: 8.215; 11: 7.265
| Std Dev. |  | 01: 0.13203; 03: 0.16654; 04: 0.10722; 06: 0.10328; 07: 0.23571; 08: 0.14634; 09: 0.20412; 10: 1.169; 11: 0.09225
| Rel.Std.Dev. |  | 01: 1.387%; 03: 1.951%; 04: 1.269%; 06: 1.27%; 07: 2.744%; 08: 1.711%; 09: 4.152%; 10: 14.06%; 11: 1.277%
| PDM3 |  | 01: 13.07%; 03: 1.407%; 04: 0.3382%; 06: -3.403%; 07: 2.021%; 08: 1.565%; 09: -41.61%; 10: -1.245%; 11: -14.19%
| Z-Score (Absolute) |  | 01: 5.1; 03: 0.43; 04: 0; 06: 1.5; 07: 0.67; 08: 0.49; 09: 16.81; 10: 0.63; 11: 5.65
| Analytical results for Pt in OREAS 905 (Indicative Value < 5 ppb) | [31 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 06: Lab |  |  | 1
| No | No | 01: 03; 03: 13; 04: 17; 06: 22 |  |  | ppb
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 06: AR*OES/MS |  |  | 0
|  |  | 01: 20g; 03: 25g; 04: 0.5g; 06: 25g |  |  | 0
| 1 | 1 | 01: < 5; 03: < 5; 04: < 2; 06: < 10 |  |  | 1
| 1 | 2 | 01: < 5; 03: < 5; 04: < 2; 06: < 10 |  |  | 1
| 1 | 3 | 01: < 5; 03: < 5; 04: < 2; 06: < 10 |  |  | 16
| 1 | 4 | 01: < 5; 03: < 5; 04: < 2; 06: < 10 |  |  | < 5
| 1 | 5 | 01: < 5; 03: < 5; 04: < 2; 06: < 10 |  |  | 20
| 1 | 6 | 01: < 5; 03: < 5; 04: < 2; 06: < 10
Mean
Median
Std Dev.
Rel.Std.Dev.
PDM3
Z-Score (Absolute) | bar: 01 to 06
| Analytical results for Rb in OREAS 905 (Certified Value 19.3 ppm) | [31 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 15 | 17 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [9 empty] | 2
| 1 | 1 | 20.6 | 27.3642 | 20.9 | 18.6 | 17.2 | 21.4 | 23.2 | 19 | 27.4 | 18.2 | 17.51 | 19.1 | 19.23 | 3.9 | 22.7 | 15.6 | 20.1 | 20.5 | 7.99 | 17.9 | 21.6 | [9 empty] | 1
| 1 | 2 | 18.6 | 26.6705 | 19.7 | 18.7 | 16.7 | 20.9 | 24.1 | 19.1 | 29 | 18.1 | 17.66 | 19.1 | 18.9 | 4.3 | 20.4 | 17.8 | 19.2 | 20 | 8.02 | 17.4 | 21.1 | [9 empty] | 0
| 1 | 3 | 20.6 | 26.57 | 20.1 | 18.8 | 17.3 | 20.2 | 24.1 | 18.8 | 26.2 | 17.9 | 18.5 | 20 | 19.17 | 4.3 | 21 | 12.6 | 19 | 20.8 | 8.17 | 18 | 21 | [9 empty] | 16
| 1 | 4 | 19.6 | 26.0594 | 20.5 | 18.8 | 17.5 | 20.7 | 23.5 | 19.7 | 26.2 | 18.1 | 18.56 | 20.3 | 18.74 | 4 | 21.4 | 21.6 | 18.1 | 21.6 | 8.44 | 18.6 | 20.2 | [9 empty] | 19.2514
| 1 | 5 | 20.4 | 27.0259 | 19.9 | 19 | 17.4 | 19.9 | 24.4 | 19.4 | 27.4 | 18 | 18.94 | 18.6 | 18.36 | 4.1 | 21.2 | 16.2 | 18.6 | 19.1 | 7.9 | 17.8 | 20 | [9 empty] | 155
| 1 | 6 | 20.8 | 26.5246 | 20.2 | 19 | 16.9 | 20.6 | 23 | 18.8 | 25.8 | 18.1 | 16.25 | 19.9 | 18.53 | 4.7 | 21.5 | 20.2 | 18.5 | 15.6 | 7.92 | 17.7 | 19.2
| Mean |  | 20.1 | 26.7024 | 20.2167 | 18.8167 | 17.1667 | 20.6167 | 23.7167 | 19.1333 | 27 | 18.0667 | 17.9033 | 19.5 | 18.8217 | 4.2167 | 21.3667 | 17.3333 | 18.9167 | 19.6 | 8.0733 | 17.9 | 20.5167
| Median |  | 20.5 | 26.6203 | 20.15 | 18.8 | 17.25 | 20.65 | 23.8 | 19.05 | 26.8 | 18.1 | 18.08 | 19.5 | 18.82 | 4.2 | 21.3 | 17 | 18.8 | 20.25 | 8.005 | 17.85 | 20.6
| Std Dev. |  | 0.84617 | 0.44861 | 0.43089 | 0.16021 | 0.30768 | 0.52694 | 0.56362 | 0.3559 | 1.1866 | 0.10328 | 0.98075 | 0.6603 | 0.34615 | 0.28577 | 0.7607 | 3.266 | 0.6969 | 2.1288 | 0.20354 | 0.4 | 0.87731
| Rel.Std.Dev. |  | 4.21% | 1.68% | 2.131% | 0.8514% | 1.792% | 2.556% | 2.376% | 1.86% | 4.395% | 0.5717% | 5.478% | 3.386% | 1.839% | 6.777% | 3.56% | 18.84% | 3.684% | 10.86% | 2.521% | 2.235% | 4.276%
| PDM3 |  | 4.408% | 38.7% | 5.014% | -2.258% | -10.83% | 7.092% | 23.19% | -0.6131% | 40.25% | -6.154% | -7.002% | 1.292% | -2.232% | -78.1% | 10.99% | -9.963% | -1.739% | 1.811% | -58.06% | -7.02% | 6.573%
| Z-Score (Absolute) |  | 0.67 | 4.03 | 0.58 | 0.17 | 1.05 | 0.79 | 2.44 | 0 | 4.19 | 0.57 | 0.65 | 0.2 | 0.17 | 7.94 | 1.19 | 0.45 | 0.12 | 0.67 | 5.93 | 0.66 | 0.74
| Analytical results for Re in OREAS 905 (Indicative Value < 0.001 ppm) | [31 empty] | Indicative
Re | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab; 16: Lab; 17: Lab; 18: Lab |  |  | 1
| No | No | 01: 01; 03: 03; 04: 06; 06: 07; 07: 08; 08: 09; 09: 11; 10: 12; 11: 13; 12: 17; 13: 22; 14: 24; 15: 28; 16: 29; 17: 30; 18: 31 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*OES/MS; 04: AR*OES/MS; 06: AR*OES/MS; 07: AR*OES/MS; 08: AR*MS; 09: AR*OES; 10: AR*MS; 11: AR*MS; 12: AR*MS; 13: AR*OES/MS; 14: AR*OES/MS; 15: AR*OES/MS; 16: AR*OES/MS; 17: AR*MS; 18: AR*OES/MS |  |  | 3
|  |  | 01: 0.4g; 03: 20g; 04: 0.5g; 06: 0.5g; 07: 20g; 08: 40g; 09: 0.15g; 10: 0.5g; 11: 25g; 12: 0.5g; 13: 25g; 14: 0.4g; 15: 0.5g; 16: 0.5g; 17: 0.4g; 18: 0.4g |  |  | 3
| 1 | 1 | 01: < 0.001; 03: < 0.001; 04: < 0.001; 06: < 0.001; 07: < 0.01; 08: < 0.01; 09: < 5; 10: < 0.001; 11: < 0.05; 12: < 0.001; 13: 0.063; 14: 0.014; 15: < 0.001; 16: < 0.001; 17: -; 18: 0.001 |  |  | 1
| 1 | 2 | 01: < 0.001; 03: < 0.001; 04: < 0.001; 06: < 0.001; 07: < 0.01; 08: < 0.01; 09: < 5; 10: < 0.001; 11: < 0.05; 12: < 0.001; 13: 0.093; 14: 0.011; 15: < 0.001; 16: < 0.001; 17: -; 18: < 0.001 |  |  | 4
| 1 | 3 | 01: < 0.001; 03: < 0.001; 04: < 0.001; 06: < 0.001; 07: < 0.01; 08: < 0.01; 09: < 5; 10: < 0.001; 11: < 0.05; 12: < 0.001; 13: 0.222; 14: 0.007; 15: < 0.001; 16: < 0.001; 17: -; 18: 0.001 |  |  | 16
| 1 | 4 | 01: < 0.001; 03: < 0.001; 04: < 0.001; 06: < 0.001; 07: < 0.01; 08: < 0.01; 09: < 5; 10: < 0.001; 11: < 0.05; 12: < 0.001; 13: < 0.005; 14: 0.006; 15: 0.001; 16: < 0.001; 17: -; 18: < 0.001 |  |  | < 0.001
| 1 | 5 | 01: < 0.001; 03: < 0.001; 04: < 0.001; 06: < 0.001; 07: < 0.01; 08: < 0.01; 09: < 5; 10: < 0.001; 11: < 0.05; 12: < 0.001; 13: 0.082; 14: 0.006; 15: < 0.001; 16: < 0.001; 17: -; 18: < 0.001 |  |  | 21
| 1 | 6 | 01: < 0.001; 03: < 0.001; 04: < 0.001; 06: < 0.001; 07: < 0.01; 08: < 0.01; 09: < 5; 10: < 0.001; 11: < 0.05; 12: < 0.001; 13: 0.057; 14: 0.012; 15: < 0.001; 16: < 0.001; 17: -; 18: < 0.001
| Mean |  | 13: 0.1034; 14: 0.0093333; 15: 0.001; 18: 0.001
| Median |  | 13: 0.082; 14: 0.009; 15: 0.001; 18: 0.001
| Std Dev. |  | 13: 0.067855; 14: 0.0034448; 18: 0
| Rel.Std.Dev. |  | 13: 65.62%; 14: 36.91%; 18: 0%
PDM3
| Z-Score (Absolute) |  | 01: NA; 03: NA; 04: NA; 06: NA; 07: NA; 08: NA; 09: NA; 10: NA; 11: NA; 12: NA; 13: 8.4; 14: 0.67; 15: 0.68; 16: NA; 17: NA; 18: 0.67
| Analytical results for Rh in OREAS 905 (Indicative Value 0.065 ppm) | [31 empty] | Indicative
Rh | Round | Replicate | 01: Lab |  |  | 1
| No | No | 01: 22 |  |  | ppm
|  |  | 01: AR*OES/MS |  |  | 3
|  |  | 01: 25g |  |  | 3
| 1 | 1 | 01: 0.06 |  |  | 1
| 1 | 2 | 01: 0.07 |  |  | 12
| 1 | 3 | 01: 0.06 |  |  | 16
| 1 | 4 | 01: 0.05 |  |  | 0.065
| 1 | 5 | 01: 0.07 |  |  | 18
| 1 | 6 | 01: 0.08
| Mean |  | 01: 0.065
| Median |  | 01: 0.065
| Std Dev. |  | 01: 0.010488
| Rel.Std.Dev. |  | 01: 16.14%
PDM3 | bar: 01
Z-Score (Absolute) | bar: 01
| Analytical results for Ru in OREAS 905 (Indicative Value 2.68 ppm) | [31 empty] | Indicative
Ru | Round | Replicate | 01: Lab |  |  | 1
| No | No | 01: 22 |  |  | ppm
|  |  | 01: AR*OES/MS |  |  | 2
|  |  | 01: 25g |  |  | 2
| 1 | 1 | 01: 2.81 |  |  | 1
| 1 | 2 | 01: 2.64 |  |  | 13
| 1 | 3 | 01: 2.66 |  |  | 16
| 1 | 4 | 01: 2.36 |  |  | 2.6783
| 1 | 5 | 01: 2.79 |  |  | 19
| 1 | 6 | 01: 2.81
| Mean |  | 01: 2.6783
| Median |  | 01: 2.725
| Std Dev. |  | 01: 0.17337
| Rel.Std.Dev. |  | 01: 6.473%
| PDM3 |  | 01: 8.882e-14%
Z-Score (Absolute) | bar: 01
| Analytical results for S in OREAS 905 (Certified Value 0.065 wt.%) | [31 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.07 | 0.055192 | 0.08 | 0.0645 | 0.04 | 0.06 | 0.052 | 0.07 | 0.06 | 0.07 | 0.06 | 0.0644 | 0.06 | 0.18 | 0.0614 | 0.06 | 0.07 | 0.0635 | 0.07 | < 0.05 | 0.06 | 0.07 | 0.07 | 0.07 | 0.065 | 0.07 | 0.063 | 0.06 |  |  | 1
| 1 | 2 | 0.07 | 0.055066 | 0.08 | 0.0665 | 0.04 | 0.06 | 0.052 | 0.07 | 0.07 | 0.07 | 0.06 | 0.0666 | 0.06 | 0.4 | 0.0616 | 0.06 | 0.067 | 0.0676 | 0.07 | 0.07 | 0.06 | 0.07 | 0.07 | 0.07 | 0.065 | 0.06 | 0.0617 | 0.06 |  |  | 24
| 1 | 3 | 0.08 | 0.05724 | 0.08 | 0.0645 | 0.04 | 0.06 | 0.053 | 0.075 | 0.06 | 0.07 | 0.06 | 0.066 | 0.06 | 0.11 | 0.0646 | 0.05 | 0.067 | 0.0682 | 0.06 | 0.07 | 0.07 | 0.07 | 0.06 | 0.07 | 0.065 | 0.07 | 0.062 | 0.06 |  |  | 16
| 1 | 4 | 0.07 | 0.056178 | 0.08 | 0.064 | 0.04 | 0.06 | 0.054 | 0.07 | 0.07 | 0.07 | 0.06 | 0.0645 | 0.06 | 0.12 | 0.062 | 0.05 | 0.072 | 0.0647 | 0.06 | 0.07 | 0.06 | 0.07 | 0.06 | 0.07 | 0.065 | 0.07 | 0.061 | 0.06 |  |  | 0.065101
| 1 | 5 | 0.07 | 0.057136 | 0.08 | 0.0625 | 0.04 | 0.06 | 0.055 | 0.075 | 0.07 | 0.07 | 0.06 | 0.0684 | 0.06 | 0.13 | 0.0612 | 0.06 | 0.076 | 0.0665 | 0.06 | 0.06 | 0.06 | 0.07 | 0.06 | 0.07 | 0.064 | 0.07 | 0.0637 | 0.06 |  |  | 156
| 1 | 6 | 0.07 | 0.055961 | 0.08 | 0.0645 | 0.04 | 0.06 | 0.053 | 0.075 | 0.06 | 0.07 | 0.06 | 0.0722 | 0.06 | 0.16 | 0.0626 | 0.06 | 0.07 | 0.0712 | 0.06 | 0.07 | 0.06 | 0.075 | 0.07 | 0.08 | 0.063 | 0.07 | 0.0608 | 0.07
| Mean |  | 0.071667 | 0.056129 | 0.08 | 0.064417 | 0.04 | 0.06 | 0.053167 | 0.0725 | 0.065 | 0.07 | 0.06 | 0.067017 | 0.06 | 0.18333 | 0.062233 | 0.056667 | 0.070333 | 0.06695 | 0.063333 | 0.068 | 0.061667 | 0.070833 | 0.065 | 0.071667 | 0.0645 | 0.068333 | 0.062033 | 0.061667
| Median |  | 0.07 | 0.05607 | 0.08 | 0.0645 | 0.04 | 0.06 | 0.053 | 0.0725 | 0.065 | 0.07 | 0.06 | 0.0663 | 0.06 | 0.145 | 0.0618 | 0.06 | 0.07 | 0.06705 | 0.06 | 0.07 | 0.06 | 0.07 | 0.065 | 0.07 | 0.065 | 0.07 | 0.06185 | 0.06
| Std Dev. |  | 0.0040825 | 0.00092593 | 0 | 0.0012813 | 0 | 0 | 0.001169 | 0.0027386 | 0.0054772 | 0 | 0 | 0.0029383 | 0 | 0.1093 | 0.0012612 | 0.005164 | 0.0033862 | 0.0027282 | 0.005164 | 0.0044721 | 0.0040825 | 0.0020412 | 0.0054772 | 0.0040825 | 0.00083666 | 0.0040825 | 0.0011325 | 0.0040825
| Rel.Std.Dev. |  | 5.696% | 1.65% | 0% | 1.989% | 0% | 0% | 2.199% | 3.777% | 8.427% | 0% | 0% | 4.384% | 0% | 59.62% | 2.027% | 9.113% | 4.815% | 4.075% | 8.154% | 6.577% | 6.62% | 2.882% | 8.427% | 5.696% | 1.297% | 5.974% | 1.826% | 6.62%
| PDM3 |  | 10.08% | -13.78% | 22.89% | -1.052% | -38.56% | -7.836% | -18.33% | 11.36% | -0.1559% | 7.524% | -7.836% | 2.942% | -7.836% | 181.6% | -4.406% | -12.96% | 8.036% | 2.839% | -2.716% | 4.452% | -5.276% | 8.804% | -0.1559% | 10.08% | -0.9239% | 4.964% | -4.713% | -5.276%
| Z-Score (Absolute) |  | 0.98 | 1.22 | 2.16 | 0.05 | 3.51 | 0.67 | 1.64 | 1.1 | 0.04 | 0.75 | 0.67 | 0.32 | 0.67 | 10.68 | 0.42 | 1.15 | 0.79 | 0.31 | 0.2 | 0.46 | 0.44 | 0.86 | 0.04 | 0.98 | 0.04 | 0.51 | 0.39 | 0.44
| Analytical results for Sb in OREAS 905 (Certified Value 1.12 ppm) | [31 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 1.18 | 1.1026 | 1.43 | 1.2 | 1.09 | 1.12 | 1.16 | 1.2 | 0.77 | < 5 | 1.43 | 0.84 | < 5 | 1.05 | < 5 | 1.21 | < 5 | < 2 | 1.8 | < 3 | 1.1 | < 3 | 0.9 | 1.12 | 1.16 | 1.2 | 1.27 | 0.92 | 0.9 |  | 1
| 1 | 2 | 1.11 | 1.0293 | 1.34 | 1.3 | 1.02 | 1.09 | 1.22 | 1.1 | 0.74 | < 5 | 1.39 | 0.85 | < 5 | 1.02 | < 5 | 1.17 | < 5 | < 2 | 1.7 | < 3 | 1.7 | < 3 | 1 | 1.11 | 1.04 | 1.1 | 1.19 | 1.06 | 0.88 |  | 0
| 1 | 3 | 1.18 | 1.0727 | 1.33 | 1.3 | 1.03 | 1.08 | 1.2 | 1.2 | 0.67 | < 5 | 1.38 | 0.77 | 5 | 1 | < 5 | 1.18 | < 5 | < 2 | 2 | < 3 | 1.5 | < 3 | 0.84 | 1.12 | 1.13 | 1.2 | 1.19 | 1.11 | 0.88 |  | 16
| 1 | 4 | 1.17 | 1.0399 | 1.39 | 1.2 | 1.07 | 1.07 | 1.19 | 1.2 | 0.59 | < 5 | 1.3 | 0.74 | < 5 | 1.04 | < 5 | 1.24 | < 5 | < 2 | 1.8 | < 3 | 1.5 | < 3 | 1.04 | 1.08 | 1.21 | 1.1 | 1.23 | 0.86 | 0.93 |  | 1.1247
| 1 | 5 | 1.15 | 1.1036 | 1.32 | 1.3 | 0.99 | 1.05 | 1.2 | 1.2 | 0.61 | < 5 | 1.47 | 0.75 | < 5 | 1.03 | < 5 | 1.2 | < 5 | < 2 | 1.7 | < 3 | 1.4 | < 3 | 0.98 | 1.08 | 1 | 1.2 | 1.24 | 0.98 | 0.78 |  | 157
| 1 | 6 | 1.19 | 1.076 | 1.36 | 1.2 | 1.02 | 1.08 | 1.17 | 1.2 | 0.56 | < 5 | 1.4 | 2.14 | < 5 | 1.02 | < 5 | 1.19 | < 5 | < 2 | 2.2 | < 3 | 1.4 | < 3 | 1.02 | 1.08 | 1.01 | 1.1 | 1.18 | 0.86 | 0.85
| Mean |  | 1.1633 | 1.0707 | 1.3617 | 1.25 | 1.0367 | 1.0817 | 1.19 | 1.1833 | 0.65667 |  | 1.395 | 1.015 | 5 | 1.0267 |  | 1.1983 |  |  | 1.8667 |  | 1.4333 |  | 0.96333 | 1.0983 | 1.0917 | 1.15 | 1.2167 | 0.965 | 0.87
| Median |  | 1.175 | 1.0744 | 1.35 | 1.25 | 1.025 | 1.08 | 1.195 | 1.2 | 0.64 |  | 1.395 | 0.805 | 5 | 1.025 |  | 1.195 |  |  | 1.8 |  | 1.45 |  | 0.99 | 1.095 | 1.085 | 1.15 | 1.21 | 0.95 | 0.88
| Std Dev. |  | 0.029439 | 0.030984 | 0.041673 | 0.054772 | 0.036697 | 0.023166 | 0.021909 | 0.040825 | 0.084774 |  | 0.056833 | 0.55306 |  | 0.017512 |  | 0.024833 |  |  | 0.19664 |  | 0.19664 |  | 0.077374 | 0.020412 | 0.087044 | 0.054772 | 0.03559 | 0.10426 | 0.051381
| Rel.Std.Dev. |  | 2.531% | 2.894% | 3.06% | 4.382% | 3.54% | 2.142% | 1.841% | 3.45% | 12.91% |  | 4.074% | 54.49% |  | 1.706% |  | 2.072% |  |  | 10.53% |  | 13.72% |  | 8.032% | 1.858% | 7.974% | 4.763% | 2.925% | 10.8% | 5.906%
| PDM3 |  | 3.44% | -4.799% | 21.07% | 11.15% | -7.823% | -3.822% | 5.811% | 5.218% | -41.61% |  | 24.04% | -9.75% | 344.6% | -8.712% |  | 6.552% |  |  | 65.98% |  | 27.45% |  | -14.34% | -2.34% | -2.933% | 2.254% | 8.182% | -14.2% | -22.64%
| Z-Score (Absolute) |  | 0.14 | 0.47 | 1.26 | 0.59 | 0.67 | 0.41 | 0.24 | 0.2 | 2.94 | NA | 1.46 | 2.14 | 22.91 | 0.73 | NA | 0.29 | NA | NA | 4.26 | NA | 1.37 | NA | 1.11 | 0.31 | 0.35 | 0 | 0.4 | 1.1 | 1.67
| Analytical results for Sc in OREAS 905 (Certified Value 1.89 ppm) | [31 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 1.9 | 1.3673 | 1.9 | 1.7 | 1.8 | 2 | 2 | 2 | 2.2 | 2 | 2.1 | 1 | < 3 | 1.9 | 1 | 1.86 | 1 | 1.8 | 0.9 | < 5 | 1.6 | < 5 | 2 | 1.9 | 2 | 1.6 | 1.8 | 2.3 | 1.6 |  | 1
| 1 | 2 | 1.8 | 1.4813 | 1.8 | 1.7 | 1.8 | 2 | 2 | 2 | 2.6 | 2 | 1.7 | 2 | < 3 | 1.9 | 1 | 1.9 | 1 | 1.8 | 1 | < 5 | 2.2 | < 5 | 2 | 1.9 | 2.1 | 1.5 | 1.7 | 2.1 | 1.5 |  | 6
| 1 | 3 | 2 | 1.6302 | 1.9 | 1.6 | 1.8 | 2 | 2 | 2 | 1.9 | 2 | 1.8 | 2 | < 3 | 2 | 1 | 1.89 | 1 | 1.7 | 1.1 | < 5 | 2.3 | < 5 | 2 | 1.8 | 2.2 | 1.6 | 1.7 | 2.1 | 1.7 |  | 16
| 1 | 4 | 1.9 | 1.4608 | 1.8 | 1.7 | 1.8 | 2 | 2 | 2 | 1.9 | 2 | 1.5 | 1 | < 3 | 2 | 1 | 1.98 | 1 | 1.8 | 1 | < 5 | 2.3 | < 5 | 2 | 1.7 | 2.2 | 1.6 | 1.8 | 2.3 | 1.8 |  | 1.8901
| 1 | 5 | 1.9 | 1.6166 | 1.8 | 1.7 | 1.8 | 1.9 | 2 | 2 | 1.9 | 2 | 1.9 | 2 | < 3 | 1.9 | 1 | 1.85 | 1 | 1.8 | 1 | < 5 | 2.3 | < 5 | 2 | 1.8 | 1.9 | 1.6 | 1.8 | 2.2 | 1.7 |  | 158
| 1 | 6 | 2.1 | 1.7315 | 1.8 | 1.8 | 1.8 | 1.9 | 2 | 2 | 1.8 | 2 | 1.9 | 1 | < 3 | 2 | 1 | 1.82 | 1 | 1.8 | 1.1 | < 5 | 2.4 | < 5 | 2 | 1.8 | 1.6 | 1.5 | 1.7 | 2.2 | 1.8
| Mean |  | 1.9333 | 1.548 | 1.8333 | 1.7 | 1.8 | 1.9667 | 2 | 2 | 2.05 | 2 | 1.8167 | 1.5 |  | 1.95 | 1 | 1.8833 | 1 | 1.7833 | 1.0167 |  | 2.1833 |  | 2 | 1.8167 | 2 | 1.5667 | 1.75 | 2.2 | 1.6833
| Median |  | 1.9 | 1.549 | 1.8 | 1.7 | 1.8 | 2 | 2 | 2 | 1.9 | 2 | 1.85 | 1.5 |  | 1.95 | 1 | 1.875 | 1 | 1.8 | 1 |  | 2.3 |  | 2 | 1.8 | 2.05 | 1.6 | 1.75 | 2.2 | 1.7
| Std Dev. |  | 0.10328 | 0.13405 | 0.05164 | 0.063246 | 0 | 0.05164 | 0 | 0 | 0.30166 | 0 | 0.20412 | 0.54772 |  | 0.054772 | 0 | 0.055377 | 0 | 0.040825 | 0.075277 |  | 0.29269 |  | 0 | 0.075277 | 0.22804 | 0.05164 | 0.054772 | 0.089443 | 0.1169
| Rel.Std.Dev. |  | 5.342% | 8.66% | 2.817% | 3.72% | 0% | 2.626% | 0% | 0% | 14.72% | 0% | 11.24% | 36.51% |  | 2.809% | 0% | 2.94% | 0% | 2.289% | 7.404% |  | 13.41% |  | 0% | 4.144% | 11.4% | 3.296% | 3.13% | 4.066% | 6.945%
| PDM3 |  | 2.29% | -18.1% | -3.001% | -10.06% | -4.765% | 4.053% | 5.817% | 5.817% | 8.462% | 5.817% | -3.883% | -20.64% |  | 3.172% | -47.09% | -0.3557% | -47.09% | -5.647% | -46.21% |  | 15.52% |  | 5.817% | -3.883% | 5.817% | -17.11% | -7.41% | 16.4% | -10.94%
| Z-Score (Absolute) |  | 0.36 | 1.48 | 0.12 | 0.75 | 0.28 | 0.52 | 0.67 | 0.67 | 0.39 | 0.67 | 0.2 | 0.67 | NA | 0.44 | 4.09 | 0.12 | 4.09 | 0.36 | 4.01 | NA | 2.1 | NA | 0.67 | 0.2 | 0.67 | 1.39 | 0.52 | 1.63 | 0.83
| Analytical results for Se in OREAS 905 (Certified Value 2.32 ppm) | [31 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  |  |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  |  |  | 2
| 1 | 1 | 2.8 | 1.73 | < 1 | 2 | 2.4 | 3.1 | 2 | 2 | 3 | < 5 | 2.3 | 2 | < 10 | 2 | 2.4 | < 10 | < 2 | 1.3 | 2.3 | 2 | 2.9 | 2.5 | < 5 | 2.5 | 2 | 2.3 |  |  |  |  | 1
| 1 | 2 | 2.4 | 2.1513 | < 1 | 2 | 2.6 | 3.1 | 2 | 2 | 3 | < 5 | 3.1 | 2 | < 10 | 2 | 2.2 | < 10 | < 2 | 1.5 | 2.5 | 2 | 2.8 | 2.4 | < 5 | 2.6 | 2 | 1.9 |  |  |  |  | 7
| 1 | 3 | 2.5 | 2.2504 | < 1 | 2 | 2.6 | 3 | 2 | 2 | 3 | < 5 | 2.8 | 2 | < 10 | 2 | 2.4 | < 10 | < 2 | 1.6 | 2.7 | 2 | 2.9 | 2.5 | < 5 | 2.7 | 2 | 2.4 |  |  |  |  | 16
| 1 | 4 | 2.5 | 2.2508 | < 1 | 2 | 2.5 | 3.2 | 2 | 2 | 3 | < 5 | 2.4 | 2 | < 10 | 2 | 2.2 | < 10 | < 2 | 1.4 | 2.7 | 2 | 2.7 | 2.7 | < 5 | 2.6 | 2 | 2.5 |  |  |  |  | 2.3169
| 1 | 5 | 2.2 | 1.9855 | < 1 | 2 | 2.2 | 2.9 | 2 | 2 | 3 | < 5 | 3.1 | 2 | < 10 | 2 | 2.3 | < 10 | < 2 | 1.5 | 2.6 | 2 | 3 | 2.3 | < 5 | 2.5 | 2 | 1.9 |  |  |  |  | 159
| 1 | 6 | 2.6 | 1.8606 | < 1 | 2 | 2.4 | 3 | 2 | 2 | 3 | < 5 | 2.9 | 3 | < 10 | 2 | 2.2 | < 10 | < 2 | 1.2 | 2.7 | 2 | 2.3 | 1.9 | < 5 | 2.6 | 2 | 2.2
| Mean |  | 2.5 | 2.0381 |  | 2 | 2.45 | 3.05 | 2 | 2 | 3 |  | 2.7667 | 2.1667 |  | 2 | 2.2833 |  |  | 1.4167 | 2.5833 | 2 | 2.7667 | 2.3833 |  | 2.5833 | 2 | 2.2
| Median |  | 2.5 | 2.0684 |  | 2 | 2.45 | 3.05 | 2 | 2 | 3 |  | 2.85 | 2 |  | 2 | 2.25 |  |  | 1.45 | 2.65 | 2 | 2.85 | 2.45 |  | 2.6 | 2 | 2.25
| Std Dev. |  | 0.2 | 0.21556 |  | 0 | 0.15166 | 0.10488 | 0 | 0 | 0 |  | 0.34448 | 0.40825 |  | 0 | 0.098319 |  |  | 0.1472 | 0.16021 | 0 | 0.25033 | 0.27142 |  | 0.075277 | 0 | 0.25298
| Rel.Std.Dev. |  | 8% | 10.58% |  | 0% | 6.19% | 3.439% | 0% | 0% | 0% |  | 12.45% | 18.84% |  | 0% | 4.306% |  |  | 10.39% | 6.202% | 0% | 9.048% | 11.39% |  | 2.914% | 0% | 11.5%
| PDM3 |  | 7.903% | -12.03% |  | -13.68% | 5.745% | 31.64% | -13.68% | -13.68% | 29.48% |  | 19.41% | -6.484% |  | -13.68% | -1.449% |  |  | -38.86% | 11.5% | -13.68% | 19.41% | 2.867% |  | 11.5% | -13.68% | -5.046%
| Z-Score (Absolute) |  | 0.72 | 0.57 | NA | 0.67 | 0.58 | 2.26 | 0.67 | 0.67 | 2.12 | NA | 1.46 | 0.21 | NA | 0.67 | 0.12 | NA | NA | 2.3 | 1.11 | 0.67 | 1.73 | 0.4 | NA | 0.95 | 0.67 | 0.12
| Analytical results for Sm in OREAS 905 (Certified Value 5.05 ppm) | [31 empty] | Certified
Sm | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  |  |  |  | 1
| No | No | 01: 03; 03: 05; 04: 09; 05: 12; 06: 13; 07: 17; 08: 22; 09: 26; 10: 30 |  |  |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*OES/MS; 09: AR*MS; 10: AR*MS |  |  |  |  | 2
|  |  | 01: 20g; 03: 0.4g; 04: 40g; 05: 0.5g; 06: 25g; 07: 0.5g; 08: 25g; 09: 40g; 10: 0.4g |  |  |  |  | 3
| 1 | 1 | 01: 5.8916; 03: 5.69; 04: 5.31; 05: 4.9; 06: 5; 07: 5.31; 08: 2.4; 09: 4.84; 10: 4.39 |  |  |  |  | 1
| 1 | 2 | 01: 5.9299; 03: 5.88; 04: 5.05; 05: 4.6; 06: 5.13; 07: 5.29; 08: 2.6; 09: 4.98; 10: 4.31 |  |  |  |  | 0
| 1 | 3 | 01: 6.0377; 03: 5.71; 04: 5.19; 05: 4.7; 06: 5.24; 07: 5.53; 08: 2.6; 09: 4.45; 10: 4.26 |  |  |  |  | 16
| 1 | 4 | 01: 5.908; 03: 5.78; 04: 5.43; 05: 4.7; 06: 4.99; 07: 5.23; 08: 2.5; 09: 5.3; 10: 4.31 |  |  |  |  | 5.0524
| 1 | 5 | 01: 5.9767; 03: 5.98; 04: 5.39; 05: 4.9; 06: 5.06; 07: 5.19; 08: 2.6; 09: 5.07; 10: 4.29 |  |  |  |  | 160
| 1 | 6 | 01: 6.0689; 03: 5.6; 04: 5.18; 05: 5; 06: 4.73; 07: 5.23; 08: 2.7; 09: 5.09; 10: 4.27
| Mean |  | 01: 5.9688; 03: 5.7733; 04: 5.2583; 05: 4.8; 06: 5.025; 07: 5.2967; 08: 2.5667; 09: 4.955; 10: 4.305
| Median |  | 01: 5.9533; 03: 5.745; 04: 5.25; 05: 4.8; 06: 5.03; 07: 5.26; 08: 2.6; 09: 5.025; 10: 4.3
| Std Dev. |  | 01: 0.072105; 03: 0.13794; 04: 0.144; 05: 0.15492; 06: 0.17167; 07: 0.12242; 08: 0.10328; 09: 0.2896; 10: 0.046368
| Rel.Std.Dev. |  | 01: 1.208%; 03: 2.389%; 04: 2.739%; 05: 3.227%; 06: 3.416%; 07: 2.311%; 08: 4.024%; 09: 5.845%; 10: 1.077%
| PDM3 |  | 01: 18.14%; 03: 14.27%; 04: 4.076%; 05: -4.995%; 06: -0.5419%; 07: 4.835%; 08: -49.2%; 09: -1.927%; 10: -14.79%
| Z-Score (Absolute) |  | 01: 2.73; 03: 2.16; 04: 0.67; 05: 0.65; 06: 0; 07: 0.65; 08: 7.1; 09: 0.2; 10: 2.08
| Analytical results for Sn in OREAS 905 (Certified Value 1.27 ppm) | [31 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [5 empty] | 2
| 1 | 1 | 1.3 | 1.4743 | 1.4 | 1.3 | 1.3 | 1.3 | 2.09 | 1 | 1.1 | < 10 | 1.14 | 1.2 | < 10 | 1.4 | 1.4 | 3 | < 5 | 1 | 1.8 | 1.2 | 1.4 | 1.4 | 4.9 | 1.4 | 1.4 | [5 empty] | 1
| 1 | 2 | 1.2 | 1.531 | 1.3 | 1.3 | 1.2 | 1.3 | 1.9 | 1 | 1.2 | < 10 | 1.18 | 1.2 | < 10 | 1.3 | 1.4 | 4 | < 5 | 1.2 | 5.1 | 1.2 | 1.4 | 1.2 | 5 | 1.3 | 1.3 | [5 empty] | 0
| 1 | 3 | 1.3 | 1.454 | 1.5 | 1.3 | 1.3 | 1.3 | 1.84 | 1 | 1.1 | < 10 | 1.23 | 1.2 | < 10 | 1.4 | 1.4 | 4 | < 5 | 1 | 7.1 | 1 | 1.4 | 1.4 | 5 | 1.3 | 1.3 | [5 empty] | 16
| 1 | 4 | 1.3 | 1.5035 | 1.4 | 1.3 | 1.3 | 1.3 | 1.98 | 1 | 1.1 | < 10 | 1.32 | 1.2 | < 10 | 1.4 | 1.5 | 4 | < 5 | 1.1 | 1.8 | 1.3 | 1.3 | 1.4 | 5 | 1.4 | 1.3 | [5 empty] | 1.2715
| 1 | 5 | 1.2 | 1.525 | 1.3 | 1.4 | 1.2 | 1.3 | 1.91 | 1 | 1.1 | < 10 | 1.41 | 1.3 | < 10 | 1.1 | 1.4 | 4 | < 5 | 1 | 1.7 | 1.2 | 1.3 | 1.2 | 5 | 1.4 | 1.3 | [5 empty] | 161
| 1 | 6 | 1.3 | 1.4719 | 1.5 | 1.3 | 1.2 | 1.3 | 2.09 | 1 | 1.1 | < 10 | 1.18 | 1.2 | < 10 | 1.4 | 1.4 | 4 | < 5 | 1 | 1.7 | 1.2 | 1.3 | 1.2 | 4.9 | 1.3 | 1.2
| Mean |  | 1.2667 | 1.4933 | 1.4 | 1.3167 | 1.25 | 1.3 | 1.9683 | 1 | 1.1167 |  | 1.2433 | 1.2167 |  | 1.3333 | 1.4167 | 3.8333 |  | 1.05 | 3.2 | 1.1833 | 1.35 | 1.3 | 4.9667 | 1.35 | 1.3
| Median |  | 1.3 | 1.4889 | 1.4 | 1.3 | 1.25 | 1.3 | 1.945 | 1 | 1.1 |  | 1.205 | 1.2 |  | 1.4 | 1.4 | 4 |  | 1 | 1.8 | 1.2 | 1.35 | 1.3 | 5 | 1.35 | 1.3
| Std Dev. |  | 0.05164 | 0.031261 | 0.089443 | 0.040825 | 0.054772 | 0 | 0.1042 | 0 | 0.040825 |  | 0.1025 | 0.040825 |  | 0.12111 | 0.040825 | 0.40825 |  | 0.083666 | 2.3341 | 0.098319 | 0.054772 | 0.10954 | 0.05164 | 0.054772 | 0.063246
| Rel.Std.Dev. |  | 4.077% | 2.093% | 6.389% | 3.101% | 4.382% | 0% | 5.294% | 0% | 3.656% |  | 8.244% | 3.355% |  | 9.083% | 2.882% | 10.65% |  | 7.968% | 72.94% | 8.309% | 4.057% | 8.427% | 1.04% | 4.057% | 4.865%
| PDM3 |  | -0.3785% | 17.45% | 10.11% | 3.554% | -1.689% | 2.243% | 54.81% | -21.35% | -12.18% |  | -2.214% | -4.311% |  | 4.865% | 11.42% | 201.5% |  | -17.42% | 151.7% | -6.933% | 6.176% | 2.243% | 290.6% | 6.176% | 2.243%
| Z-Score (Absolute) |  | 0.31 | 1.36 | 0.67 | 0.06 | 0.43 | 0.06 | 4.86 | 2.27 | 1.41 | NA | 0.48 | 0.67 | NA | 0.18 | 0.8 | 18.57 | NA | 1.9 | 8.18 | 0.92 | 0.31 | 0.06 | 26.91 | 0.31 | 0.06
| Analytical results for Sr in OREAS 905 (Certified Value 12.3 ppm) | [31 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 12.3 | 13.3634 | 13.3 | 12.6 | 11.4 | 13 | 11.8 | 13 | 11.2 | 13 | 11.7 | 10 | 13.4 | 12 | 13.28 | 10.1 | 13.1 | 10 | 11 | 12 | 12 | 12.1 | 15 | 14.5 | 15 | 12.1 | 12.7 | 11.9 |  |  | 1
| 1 | 2 | 10.4 | 12.6434 | 12.8 | 12.9 | 11.2 | 12.9 | 11.9 | 12 | 11.9 | 12 | 11.8 | 10 | 12.7 | 12 | 12.88 | 9.9 | 12.8 | 12 | 11 | 13 | 12 | 12.3 | 14.3 | 14.2 | 14 | 11.7 | 12 | 11.8 |  |  | 0
| 1 | 3 | 12.1 | 12.9584 | 12.6 | 12.8 | 11.4 | 13.1 | 11.5 | 13 | 11.5 | 13 | 12 | 10 | 13.2 | 12 | 13.31 | 9.5 | 12.8 | 12 | 11 | 13 | 12 | 8.5 | 14 | 14.9 | 14.6 | 12.2 | 12 | 12.3 |  |  | 16
| 1 | 4 | 12.1 | 12.674 | 12.7 | 10.8 | 11.7 | 13 | 11.4 | 12 | 11.7 | 13 | 12.1 | 10 | 13.2 | 11 | 13.22 | 9.9 | 13.6 | 11 | 11 | 13 | 12 | 14.8 | 13.5 | 15.5 | 14.4 | 12.6 | 11.8 | 12.5 |  |  | 12.3315
| 1 | 5 | 11.8 | 13.5374 | 12.5 | 12.3 | 11.2 | 12.9 | 11.6 | 13 | 11.9 | 12 | 12.4 | 10 | 13.2 | 12 | 12.92 | 9.9 | 13.6 | 11 | 11 | 14 | 12 | 12.4 | 14 | 13.7 | 15 | 11.9 | 12.2 | 11.8 |  |  | 162
| 1 | 6 | 12.1 | 13.0272 | 12.7 | 10.8 | 11.3 | 12.8 | 11.8 | 13 | 11.6 | 13 | 11.5 | 10 | 12.7 | 12 | 12.99 | 10.6 | 13.3 | 12 | 11 | 14 | 12 | 13.8 | 13.6 | 11.7 | 14.2 | 12 | 11.8 | 12.4
| Mean |  | 11.8 | 13.034 | 12.7667 | 12.0333 | 11.3667 | 12.95 | 11.6667 | 12.6667 | 11.6333 | 12.6667 | 11.9167 | 10 | 13.0667 | 11.8333 | 13.1 | 9.9833 | 13.2 | 11.3333 | 11 | 13.1667 | 12 | 12.3167 | 14.0667 | 14.0833 | 14.5333 | 12.0833 | 12.0833 | 12.1167
| Median |  | 12.1 | 12.9928 | 12.7 | 12.45 | 11.35 | 12.95 | 11.7 | 13 | 11.65 | 13 | 11.9 | 10 | 13.2 | 12 | 13.105 | 9.9 | 13.2 | 11.5 | 11 | 13 | 12 | 12.35 | 14 | 14.35 | 14.5 | 12.05 | 12 | 12.1
| Std Dev. |  | 0.70427 | 0.36051 | 0.28048 | 0.97707 | 0.18619 | 0.10488 | 0.19664 | 0.5164 | 0.26583 | 0.5164 | 0.31885 | 0 | 0.29439 | 0.40825 | 0.19173 | 0.36009 | 0.36332 | 0.8165 | 0 | 0.75277 | 0 | 2.1442 | 0.54283 | 1.3182 | 0.41312 | 0.30605 | 0.33714 | 0.31885
| Rel.Std.Dev. |  | 5.968% | 2.766% | 2.197% | 8.12% | 1.638% | 0.8099% | 1.685% | 4.077% | 2.285% | 4.077% | 2.676% | 0% | 2.253% | 3.45% | 1.464% | 3.607% | 2.752% | 7.204% | 0% | 5.717% | 0% | 17.41% | 3.859% | 9.36% | 2.843% | 2.533% | 2.79% | 2.632%
| PDM3 |  | -4.31% | 5.697% | 3.529% | -2.418% | -7.824% | 5.016% | -5.391% | 2.718% | -5.661% | 2.718% | -3.364% | -18.91% | 5.962% | -4.039% | 6.232% | -19.04% | 7.043% | -8.094% | -10.8% | 6.773% | -2.688% | -0.1199% | 14.07% | 14.21% | 17.86% | -2.012% | -2.012% | -1.742%
| Z-Score (Absolute) |  | 0.02 | 0.84 | 0.5 | 0.06 | 0.66 | 0.76 | 0.39 | 0.51 | 0.42 | 0.51 | 0.16 | 1.89 | 0.87 | 0.24 | 0.9 | 2.01 | 0.99 | 0.69 | 0.99 | 0.96 | 0.09 | 0.88 | 1.77 | 1.78 | 2.19 | 0.01 | 0.01 | 0.01
| Analytical results for Ta in OREAS 905 (Indicative Value < 0.01 ppm) | [31 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [10 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 13 | 15 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [10 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | [10 empty] | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 0.25g | 0.5g | 25g | 0.25g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [10 empty] | 3
| 1 | 1 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.09 | < 0.05 | < 0.05 | < 0.05 | < 5 | 0.7 | 0.2 | < 0.005 | < 0.01 | < 0.01 | < 0.05 | 0.01 | 0.08 | [10 empty] | 1
| 1 | 2 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.09 | < 0.05 | < 0.05 | < 0.05 | < 5 | 1 | < 0.1 | < 0.005 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.1 | [10 empty] | 14
| 1 | 3 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.07 | < 0.05 | < 0.05 | < 0.05 | < 5 | 1 | < 0.1 | < 0.005 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.07 | [10 empty] | 16
| 1 | 4 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | 0.06 | < 0.05 | < 0.05 | < 0.05 | < 5 | 1.1 | < 0.1 | < 0.005 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.07 | [10 empty] | < 0.01
| 1 | 5 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 5 | 1.1 | < 0.1 | < 0.005 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.06 | [10 empty] | 20
| 1 | 6 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.01 | < 0.01 | < 0.01 | < 0.05 | < 0.05 | < 0.05 | < 0.05 | < 5 | 0.7 | < 0.1 | < 0.005 | < 0.01 | < 0.01 | < 0.05 | < 0.01 | 0.07
| Mean |  | 09: 0.0775; 15: 0.93333; 16: 0.2; 23: 0.01; 24: 0.075
| Median |  | 09: 0.08; 15: 1; 16: 0.2; 23: 0.01; 24: 0.07
| Std Dev. |  | 09: 0.015; 15: 0.18619; 24: 0.013784
| Rel.Std.Dev. |  | 09: 19.35%; 15: 19.95%; 24: 18.38%
PDM3
| Z-Score (Absolute) |  | NA | NA | NA | NA | NA | NA | NA | 0 | NA | NA | NA | NA | 58.89 | 0.67 | NA | NA | NA | NA | 3.65 | 0.67
| Analytical results for Tb in OREAS 905 (Certified Value 0.44 ppm) | [31 empty] | Certified
Tb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 12: Lab; 13: Lab; 14: Lab; 15: Lab |  |  | 1
| No | No | 01: 03; 03: 04; 04: 05; 05: 09; 06: 10; 07: 12; 08: 13; 09: 15; 10: 17; 12: 22; 13: 26; 14: 30; 15: 32 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*MS; 12: AR*OES/MS; 13: AR*MS; 14: AR*MS; 15: AR*OES/MS |  |  | 2
|  |  | 01: 20g; 03: 0.25g; 04: 0.4g; 05: 40g; 06: 0.25g; 07: 0.5g; 08: 25g; 09: 0.25g; 10: 0.5g; 12: 25g; 13: 40g; 14: 0.4g; 15: 0.25g |  |  | 2
| 1 | 1 | 01: 0.49595; 03: 0.46; 04: 0.48; 05: 0.375; 06: 0.48; 07: 0.4; 08: 0.38; 09: 0.5; 10: 0.454; 12: < 0.2; 13: 0.31; 14: 1.77; 15: 0.5 |  |  | 1
| 1 | 2 | 01: 0.45902; 03: 0.45; 04: 0.48; 05: 0.36; 06: 0.49; 07: 0.4; 08: 0.4; 09: 0.5; 10: 0.436; 12: < 0.2; 13: 0.38; 14: 1.74; 15: 0.47 |  |  | 0
| 1 | 3 | 01: 0.49954; 03: 0.45; 04: 0.49; 05: 0.38; 06: 0.47; 07: 0.4; 08: 0.4; 09: 0.51; 10: 0.457; 12: < 0.2; 13: 0.345; 14: 1.78; 15: 0.47 |  |  | 16
| 1 | 4 | 01: 0.49537; 03: 0.44; 04: 0.48; 05: 0.37; 06: 0.48; 07: 0.4; 08: 0.39; 09: 0.51; 10: 0.437; 12: < 0.2; 13: 0.39; 14: 1.75; 15: 0.46 |  |  | 0.43862
| 1 | 5 | 01: 0.48299; 03: 0.44; 04: 0.5; 05: 0.355; 06: 0.49; 07: 0.4; 08: 0.39; 09: 0.49; 10: 0.44; 12: < 0.2; 13: 0.335; 14: 1.75; 15: 0.44 |  |  | 163
| 1 | 6 | 01: 0.49191; 03: 0.45; 04: 0.47; 05: 0.355; 06: 0.48; 07: 0.4; 08: 0.36; 09: 0.5; 10: 0.446; 12: < 0.2; 13: 0.385; 14: 1.75; 15: 0.43
| Mean |  | 01: 0.48746; 03: 0.44833; 04: 0.48333; 05: 0.36583; 06: 0.48167; 07: 0.4; 08: 0.38667; 09: 0.50167; 10: 0.445; 13: 0.3575; 14: 1.7567; 15: 0.46167
| Median |  | 01: 0.49364; 03: 0.45; 04: 0.48; 05: 0.365; 06: 0.48; 07: 0.4; 08: 0.39; 09: 0.5; 10: 0.443; 13: 0.3625; 14: 1.75; 15: 0.465
| Std Dev. |  | 01: 0.015027; 03: 0.0075277; 04: 0.010328; 05: 0.010685; 06: 0.0075277; 07: 6.0809e-17; 08: 0.015055; 09: 0.0075277; 10: 0.0088994; 13: 0.032365; 14: 0.015055; 15: 0.024833
| Rel.Std.Dev. |  | 01: 3.083%; 03: 1.679%; 04: 2.137%; 05: 2.921%; 06: 1.563%; 07: 1.52e-14%; 08: 3.894%; 09: 1.501%; 10: 2%; 13: 9.053%; 14: 0.857%; 15: 5.379%
| PDM3 |  | 01: 11.14%; 03: 2.215%; 04: 10.19%; 05: -16.59%; 06: 9.814%; 07: -8.805%; 08: -11.84%; 09: 14.37%; 10: 1.455%; 13: -18.49%; 14: 300.5%; 15: 5.254%
| Z-Score (Absolute) |  | 01: 0.61; 03: 0.11; 04: 0.45; 05: 1.42; 06: 0.42; 07: 0.87; 08: 1.09; 09: 0.74; 10: 0.16; 12: NA; 13: 1.55; 14: 20.7; 15: 0.11
| Analytical results for Te in OREAS 905 (Certified Value 0.065 ppm) | [31 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 17 | 19 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | [7 empty] | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 0.25g | 0.5g | 0.5g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g | [7 empty] | 3
| 1 | 1 | 0.08 | 0.061823 | < 0.05 | < 0.1 | 0.07 | 0.05 | 0.03 | < 0.2 | 0.14 | < 5 | < 0.02 | 0.06 | 0.07 | 0.07 | < 2 | < 0.1 | 0.06 | 0.07 | 0.06 | < 0.1 | 0.06 | 0.06 | 0.08 | [7 empty] | 1
| 1 | 2 | 0.07 | 0.041428 | < 0.05 | < 0.1 | 0.07 | 0.07 | 0.04 | < 0.2 | 0.09 | < 5 | < 0.02 | 0.05 | 0.08 | 0.06 | < 2 | < 0.1 | 0.06 | 0.1 | 0.06 | < 0.1 | 0.06 | 0.06 | 0.1 | [7 empty] | 8
| 1 | 3 | 0.06 | 0.045939 | < 0.05 | < 0.1 | 0.06 | 0.07 | 0.04 | < 0.2 | 0.07 | < 5 | < 0.02 | 0.07 | 0.07 | 0.07 | < 2 | 0.1 | 0.04 | 0.09 | 0.07 | < 0.1 | 0.07 | 0.06 | 0.08 | [7 empty] | 16
| 1 | 4 | 0.06 | 0.057056 | < 0.05 | < 0.1 | 0.08 | 0.06 | 0.03 | < 0.2 | 0.05 | < 5 | < 0.02 | 0.06 | 0.08 | 0.06 | < 2 | < 0.1 | 0.06 | 0.07 | 0.07 | < 0.1 | 0.07 | 0.05 | 0.11 | [7 empty] | 0.065111
| 1 | 5 | 0.06 | 0.062021 | < 0.05 | < 0.1 | 0.07 | 0.08 | 0.03 | < 0.2 | < 0.05 | < 5 | < 0.02 | 0.07 | 0.06 | 0.07 | < 2 | < 0.1 | 0.06 | 0.08 | 0.06 | < 0.1 | 0.07 | 0.08 | 0.03 | [7 empty] | 164
| 1 | 6 | 0.07 | 0.05373 | < 0.05 | < 0.1 | 0.06 | 0.07 | 0.04 | < 0.2 | < 0.05 | < 5 | < 0.02 | 2.34 | 0.07 | 0.07 | < 2 | 0.1 | 0.06 | 0.11 | 0.05 | < 0.1 | 0.07 | 0.09 | 0.08
| Mean |  | 01: 0.066667; 03: 0.053666; 06: 0.068333; 07: 0.066667; 08: 0.035; 10: 0.0875; 14: 0.44167; 15: 0.071667; 16: 0.066667; 18: 0.1; 19: 0.056667; 22: 0.086667; 23: 0.061667; 25: 0.066667; 26: 0.066667; 28: 0.08
| Median |  | 01: 0.065; 03: 0.055393; 06: 0.07; 07: 0.07; 08: 0.035; 10: 0.08; 14: 0.065; 15: 0.07; 16: 0.07; 18: 0.1; 19: 0.06; 22: 0.085; 23: 0.06; 25: 0.07; 26: 0.06; 28: 0.08
| Std Dev. |  | 01: 0.008165; 03: 0.0084538; 06: 0.0075277; 07: 0.010328; 08: 0.0054772; 10: 0.038622; 14: 0.93002; 15: 0.0075277; 16: 0.005164; 18: 0; 19: 0.008165; 22: 0.01633; 23: 0.0075277; 25: 0.005164; 26: 0.015055; 28: 0.027568
| Rel.Std.Dev. |  | 01: 12.25%; 03: 15.75%; 06: 11.02%; 07: 15.49%; 08: 15.65%; 10: 44.14%; 14: 210.6%; 15: 10.5%; 16: 7.746%; 18: 0%; 19: 14.41%; 22: 18.84%; 23: 12.21%; 25: 7.746%; 26: 22.58%; 28: 34.46%
| PDM3 |  | 01: 2.389%; 03: -17.58%; 06: 4.949%; 07: 2.389%; 08: -46.25%; 10: 34.39%; 14: 578.3%; 15: 10.07%; 16: 2.389%; 18: 53.58%; 19: -12.97%; 22: 33.11%; 23: -5.29%; 25: 2.389%; 26: 2.389%; 28: 22.87%
| Z-Score (Absolute) |  | 0 | 1.75 | NA | NA | 0.22 | 0 | 4.27 | NA | 2.81 | NA | NA | 0.63 | 0.67 | 0 | NA | 4.5 | 1.35 | 2.7 | 0.67 | NA | 0 | 0 | 0.99
| Analytical results for Th in OREAS 905 (Certified Value 8.88 ppm) | [31 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 14 | 15 | 17 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.5g | 25g | 10g | 0.25g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g | [5 empty] | 3
| 1 | 1 | 9 | 9.2586 | 9.8 | 8.71 | 8.3 | 9.4 | 8.66 | 8.5 | 9.8 | 7.7 | 7.59 | < 10 | 9.3 | 9.09 | 1.5 | 11 | 9.3 | 8 | 7.98 | 9.9 | 10.9 | < 10 | 8 | 9.2 | 9.1 | [5 empty] | 1
| 1 | 2 | 7.6 | 8.9896 | 9.4 | 8.9 | 8.1 | 9.2 | 8.78 | 8.45 | 9.9 | 7.7 | 8.02 | < 10 | 9.2 | 8.89 | 1.6 | 11 | 10.2 | 9 | 8.46 | 10.1 | 10.7 | < 10 | 8.3 | 9.1 | 9 | [5 empty] | 0
| 1 | 3 | 9.1 | 9.1252 | 9.6 | 8.8 | 8.4 | 9.6 | 8.69 | 8.7 | 9.5 | 7.8 | 8.29 | < 10 | 9.2 | 9.23 | 1.6 | 11 | 10.2 | 9 | 7.3 | 9.9 | 11.1 | < 10 | 8.5 | 9.5 | 7.7 | [5 empty] | 16
| 1 | 4 | 8.5 | 9.1221 | 9.6 | 8.84 | 8.5 | 9.4 | 8.72 | 8.65 | 9.7 | 7.9 | 8 | < 10 | 9.3 | 9.01 | 1.8 | 11 | 10.3 | 9 | 8.62 | 9.8 | 11.5 | < 10 | 8.4 | 9.3 | 8.7 | [5 empty] | 8.8816
| 1 | 5 | 8.9 | 9.2163 | 9.6 | 9.05 | 8 | 9.2 | 8.69 | 8.65 | 10.1 | 7.9 | 8.08 | < 10 | 8.3 | 8.92 | 1.6 | 12 | 10.3 | 8 | 8.32 | 9.5 | 9.9 | < 10 | 8.3 | 9.1 | 7.8 | [5 empty] | 165
| 1 | 6 | 8.8 | 9.3162 | 9.8 | 8.61 | 8.2 | 9.4 | 8.61 | 8.4 | 9.6 | 7.8 | 7.35 | < 10 | 8.9 | 9.13 | 1.8 | 11 | 10.4 | 8 | 8.2 | 9.8 | 7.8 | < 10 | 8.5 | 9.1 | 7.8
| Mean |  | 8.65 | 9.1713 | 9.6333 | 8.8183 | 8.25 | 9.3667 | 8.6917 | 8.5583 | 9.7667 | 7.8 | 7.8883 |  | 9.0333 | 9.045 | 1.65 | 11.1667 | 10.1167 | 8.5 | 8.1467 | 9.8333 | 10.3167 |  | 8.3333 | 9.2167 | 8.35
| Median |  | 8.85 | 9.1708 | 9.6 | 8.82 | 8.25 | 9.4 | 8.69 | 8.575 | 9.75 | 7.8 | 8.01 |  | 9.2 | 9.05 | 1.6 | 11 | 10.25 | 8.5 | 8.26 | 9.85 | 10.8 |  | 8.35 | 9.15 | 8.25
| Std Dev. |  | 0.55408 | 0.11679 | 0.15055 | 0.1525 | 0.18708 | 0.15055 | 0.057067 | 0.12416 | 0.21602 | 0.089443 | 0.34833 |  | 0.38816 | 0.12988 | 0.12247 | 0.40825 | 0.40702 | 0.54772 | 0.46898 | 0.19664 | 1.3423 |  | 0.18619 | 0.16021 | 0.65345
| Rel.Std.Dev. |  | 6.406% | 1.273% | 1.563% | 1.729% | 2.268% | 1.607% | 0.6566% | 1.451% | 2.212% | 1.147% | 4.416% |  | 4.297% | 1.436% | 7.423% | 3.656% | 4.023% | 6.444% | 5.757% | 2% | 13.01% |  | 2.234% | 1.738% | 7.826%
| PDM3 |  | -2.608% | 3.262% | 8.463% | -0.7129% | -7.112% | 5.461% | -2.139% | -3.64% | 9.965% | -12.18% | -11.18% |  | 1.708% | 1.839% | -81.42% | 25.73% | 13.91% | -4.297% | -8.275% | 10.72% | 16.16% |  | -6.174% | 3.772% | -5.986%
| Z-Score (Absolute) |  | 0 | 0.4 | 0.99 | 0.05 | 0.78 | 0.65 | 0.22 | 0.39 | 1.16 | 1.36 | 1.24 | NA | 0.41 | 0.24 | 9.23 | 2.95 | 1.82 | 0.46 | 0.91 | 1.25 | 2.51 | NA | 0.67 | 0.38 | 0.65
| Analytical results for Ti in OREAS 905 (Certified Value 0.019 wt.%) | [31 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | wt.%
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 3
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 3
| 1 | 1 | 0.022 | 0.028482 | 0.02 | 0.02 | 0.02 | 0.023 | 0.0287 | 0.02 | 0.02 | 0.02 | 0.0135 | 0.02 | 0.03 | 0.02 | 0.0196 | 0.02 | 0.02 | 0.0097 | 0.015 | 0.012 | 0.017 | 0.01 | 0.021 | 0.019 | 0.017 | 0.02 | 0.0254 | 0.018 |  |  | 1
| 1 | 2 | 0.021 | 0.029152 | 0.02 | 0.021 | 0.02 | 0.023 | 0.0285 | 0.02 | 0.02 | 0.02 | 0.0141 | 0.02 | 0.02 | 0.01 | 0.0188 | 0.02 | 0.02 | 0.0102 | 0.016 | 0.012 | 0.017 | 0.01 | 0.02 | 0.019 | 0.017 | 0.02 | 0.0241 | 0.018 |  |  | 0
| 1 | 3 | 0.022 | 0.028439 | 0.02 | 0.019 | 0.021 | 0.022 | 0.029 | 0.02 | 0.02 | 0.02 | 0.0134 | 0.02 | 0.03 | 0.02 | 0.0198 | 0.02 | 0.02 | 0.0099 | 0.016 | 0.012 | 0.017 | 0.01 | 0.021 | 0.019 | 0.018 | 0.019 | 0.0243 | 0.019 |  |  | 16
| 1 | 4 | 0.022 | 0.027517 | 0.02 | 0.02 | 0.02 | 0.022 | 0.0297 | 0.02 | 0.02 | 0.02 | 0.0134 | 0.02 | 0.03 | 0.02 | 0.0194 | 0.02 | 0.02 | 0.01 | 0.016 | 0.012 | 0.017 | 0.01 | 0.021 | 0.019 | 0.019 | 0.02 | 0.0244 | 0.019 |  |  | 0.019173
| 1 | 5 | 0.022 | 0.029703 | 0.02 | 0.0205 | 0.021 | 0.022 | 0.0299 | 0.02 | 0.02 | 0.02 | 0.0135 | 0.02 | 0.03 | 0.02 | 0.0189 | 0.02 | 0.02 | 0.0097 | 0.016 | 0.012 | 0.017 | 0.01 | 0.02 | 0.018 | 0.016 | 0.02 | 0.0242 | 0.019 |  |  | 166
| 1 | 6 | 0.022 | 0.029662 | 0.02 | 0.021 | 0.02 | 0.022 | 0.0292 | 0.02 | 0.02 | 0.02 | 0.0119 | 0.02 | 0.02 | 0.02 | 0.0197 | 0.02 | 0.02 | 0.0113 | 0.015 | 0.012 | 0.017 | 0.01 | 0.021 | 0.018 | 0.017 | 0.019 | 0.0234 | 0.02
| Mean |  | 0.021833 | 0.028826 | 0.02 | 0.02025 | 0.020333 | 0.022333 | 0.029167 | 0.02 | 0.02 | 0.02 | 0.0133 | 0.02 | 0.026667 | 0.018333 | 0.019367 | 0.02 | 0.02 | 0.010133 | 0.015667 | 0.012 | 0.017 | 0.01 | 0.020667 | 0.018667 | 0.017333 | 0.019667 | 0.0243 | 0.018833
| Median |  | 0.022 | 0.028817 | 0.02 | 0.02025 | 0.02 | 0.022 | 0.0291 | 0.02 | 0.02 | 0.02 | 0.01345 | 0.02 | 0.03 | 0.02 | 0.0195 | 0.02 | 0.02 | 0.00995 | 0.016 | 0.012 | 0.017 | 0.01 | 0.021 | 0.019 | 0.017 | 0.02 | 0.02425 | 0.019
| Std Dev. |  | 0.00040825 | 0.00084326 | 0 | 0.00075829 | 0.0005164 | 0.0005164 | 0.00055015 | 0 | 0 | 0 | 0.00073485 | 0 | 0.005164 | 0.0040825 | 0.00042269 | 0 | 0 | 0.00060222 | 0.0005164 | 1.9003e-18 | 0 | 0 | 0.0005164 | 0.0005164 | 0.0010328 | 0.0005164 | 0.00064498 | 0.00075277
| Rel.Std.Dev. |  | 1.87% | 2.925% | 0% | 3.745% | 2.54% | 2.312% | 1.886% | 0% | 0% | 0% | 5.525% | 0% | 19.36% | 22.27% | 2.183% | 0% | 0% | 5.943% | 3.296% | 1.584e-14% | 0% | 0% | 2.499% | 2.766% | 5.958% | 2.626% | 2.654% | 3.997%
| PDM3 |  | 13.87% | 50.34% | 4.311% | 5.615% | 6.049% | 16.48% | 52.12% | 4.311% | 4.311% | 4.311% | -30.63% | 4.311% | 39.08% | -4.382% | 1.008% | 4.311% | 4.311% | -47.15% | -18.29% | -37.41% | -11.34% | -47.84% | 7.788% | -2.643% | -9.597% | 2.572% | 26.74% | -1.774%
| Z-Score (Absolute) |  | 0.99 | 4.76 | 0 | 0.13 | 0.18 | 1.26 | 4.94 | 0 | 0 | 0 | 3.46 | 0 | 0 | 0.9 | 0.34 | 0 | 0 | 5.45 | 2.34 | 4.32 | 1.62 | 5.39 | 0.36 | 0.72 | 1.44 | 0.18 | 2.32 | 0.63
| Analytical results for Tl in OREAS 905 (Certified Value 0.11 ppm) | [31 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  |  | 2
| 1 | 1 | 0.1 | 0.15 | 0.11 | 0.1 | 0.11 | 0.12 | 0.11 | < 0.1 | 0.12 | < 5 | 0.22 | 0.1 | < 10 | 0.12 | 0.11 | < 10 | < 10 | 0.08 | < 5 | 0.3 | < 5 | 0.1 | 0.1 | 0.13 | < 0.1 | 0.1 | 0.15 | 0.11 |  |  | 1
| 1 | 2 | 0.1 | 0.15317 | 0.11 | 0.1 | 0.1 | 0.11 | 0.12 | < 0.1 | 0.13 | < 5 | 0.16 | 0.11 | < 10 | 0.12 | 0.11 | < 10 | < 10 | 0.14 | < 5 | 0.3 | < 5 | 0.1 | 0.1 | 0.13 | < 0.1 | 0.1 | 0.15 | 0.1 |  |  | 0
| 1 | 3 | 0.12 | 0.13912 | 0.12 | 0.1 | 0.11 | 0.12 | 0.13 | 0.1 | 0.11 | < 5 | 0.12 | 0.11 | < 10 | 0.13 | 0.12 | < 10 | < 10 | 0.17 | < 5 | 0.2 | < 5 | 0.11 | 0.11 | 0.13 | < 0.1 | 0.11 | 0.15 | 0.1 |  |  | 16
| 1 | 4 | 0.11 | 0.14557 | 0.12 | 0.1 | 0.1 | 0.12 | 0.12 | < 0.1 | 0.12 | < 5 | 0.11 | 0.11 | < 10 | 0.13 | 0.11 | < 10 | < 10 | 0.12 | < 5 | 0.2 | < 5 | 0.1 | 0.11 | 0.13 | < 0.1 | 0.11 | 0.14 | 0.1 |  |  | 0.10972
| 1 | 5 | 0.11 | 0.13312 | 0.12 | 0.1 | 0.09 | 0.11 | 0.12 | 0.1 | 0.12 | < 5 | 0.11 | 0.11 | < 10 | 0.11 | 0.11 | < 10 | < 10 | 0.09 | < 5 | 0.2 | < 5 | 0.11 | 0.1 | 0.11 | < 0.1 | 0.11 | 0.14 | 0.08 |  |  | 167
| 1 | 6 | 0.11 | 0.16032 | 0.12 | 0.1 | 0.1 | 0.11 | 0.11 | < 0.1 | 0.11 | < 5 | 0.12 | 0.16 | < 10 | 0.12 | 0.11 | < 10 | < 10 | 0.08 | < 5 | 0.2 | < 5 | 0.11 | 0.11 | 0.12 | < 0.1 | 0.11 | 0.14 | 0.09
| Mean |  | 0.10833 | 0.14688 | 0.11667 | 0.1 | 0.10167 | 0.115 | 0.11833 | 0.1 | 0.11833 |  | 0.14 | 0.11667 |  | 0.12167 | 0.11167 |  |  | 0.11333 |  | 0.23333 |  | 0.105 | 0.105 | 0.125 |  | 0.10667 | 0.145 | 0.096667
| Median |  | 0.11 | 0.14778 | 0.12 | 0.1 | 0.1 | 0.115 | 0.12 | 0.1 | 0.12 |  | 0.12 | 0.11 |  | 0.12 | 0.11 |  |  | 0.105 |  | 0.2 |  | 0.105 | 0.105 | 0.13 |  | 0.11 | 0.145 | 0.1
| Std Dev. |  | 0.0075277 | 0.009811 | 0.005164 | 1.5202e-17 | 0.0075277 | 0.0054772 | 0.0075277 | 0 | 0.0075277 |  | 0.043359 | 0.021602 |  | 0.0075277 | 0.0040825 |  |  | 0.036697 |  | 0.05164 |  | 0.0054772 | 0.0054772 | 0.0083666 |  | 0.005164 | 0.0054772 | 0.010328
| Rel.Std.Dev. |  | 6.949% | 6.679% | 4.426% | 1.52e-14% | 7.404% | 4.763% | 6.361% | 0% | 6.361% |  | 30.97% | 18.52% |  | 6.187% | 3.656% |  |  | 32.38% |  | 22.13% |  | 5.216% | 5.216% | 6.693% |  | 4.841% | 3.777% | 10.68%
| PDM3 |  | -1.266% | 33.87% | 6.329% | -8.861% | -7.342% | 4.81% | 7.848% | -8.861% | 7.848% |  | 27.59% | 6.329% |  | 10.89% | 1.772% |  |  | 3.291% |  | 112.7% |  | -4.304% | -4.304% | 13.92% |  | -2.785% | 32.15% | -11.9%
| Z-Score (Absolute) |  | 0.34 | 3.56 | 0.51 | 1.18 | 1.01 | 0.34 | 0.67 | 1.18 | 0.67 | NA | 0.34 | 0.37 | NA | 1.01 | 0 | NA | NA | 0.98 | NA | 12.31 | NA | 0.67 | 0.67 | 1.35 | NA | 0.51 | 3.37 | 1.52
| Analytical results for Tm in OREAS 905 (Certified Value 0.047 ppm) | [31 empty] | Certified
Tm | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  | 1
| No | No | 01: 03; 03: 09; 04: 12; 05: 13; 06: 17; 07: 22; 08: 26 |  |  | ppm
|  |  | 01: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*OES/MS; 08: AR*MS |  |  | 3
|  |  | 01: 20g; 03: 40g; 04: 0.5g; 05: 25g; 06: 0.5g; 07: 25g; 08: 40g |  |  | 3
| 1 | 1 | 01: 0.056185; 03: 0.05; 04: < 0.1; 05: 0.04; 06: 0.05; 07: < 0.01; 08: 0.04 |  |  | 1
| 1 | 2 | 01: 0.060624; 03: 0.045; 04: < 0.1; 05: 0.04; 06: 0.05; 07: < 0.01; 08: 0.045 |  |  | 0
| 1 | 3 | 01: 0.056737; 03: 0.05; 04: < 0.1; 05: 0.04; 06: 0.05; 07: < 0.01; 08: 0.04 |  |  | 16
| 1 | 4 | 01: 0.05425; 03: 0.05; 04: < 0.1; 05: 0.04; 06: 0.04; 07: < 0.01; 08: 0.05 |  |  | 0.0473
| 1 | 5 | 01: 0.063687; 03: 0.045; 04: < 0.1; 05: 0.04; 06: 0.05; 07: < 0.01; 08: 0.04 |  |  | 168
| 1 | 6 | 01: 0.062525; 03: 0.045; 04: < 0.1; 05: 0.03; 06: 0.05; 07: < 0.01; 08: 0.045
| Mean |  | 01: 0.059001; 03: 0.0475; 05: 0.038333; 06: 0.048333; 08: 0.043333
| Median |  | 01: 0.05868; 03: 0.0475; 05: 0.04; 06: 0.05; 08: 0.0425
| Std Dev. |  | 01: 0.0038115; 03: 0.0027386; 05: 0.0040825; 06: 0.0040825; 08: 0.0040825
| Rel.Std.Dev. |  | 01: 6.46%; 03: 5.766%; 05: 10.65%; 06: 8.447%; 08: 9.421%
| PDM3 |  | 01: 24.74%; 03: 0.4223%; 05: -18.96%; 06: 2.184%; 08: -8.387%
| Z-Score (Absolute) |  | 01: 1.86; 03: 0; 04: NA; 05: 1.48; 06: 0.13; 07: NA; 08: 0.67
| Analytical results for U in OREAS 905 (Certified Value 2.37 ppm) | [31 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | [6 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g | [6 empty] | 3
| 1 | 1 | 2.31 | 2.6568 | 2.51 | 2.4 | 2.18 | 2.3 | 2.54 | 2.2 | 2.31 | < 10 | 2.5 | 2.04 | < 10 | 2.55 | 2.4 | < 0.1 | 2.5 | 2 | 2.59 | 2.57 | < 10 | 2.05 | 2.5 | 2.4 | [6 empty] | 1
| 1 | 2 | 2.17 | 2.6282 | 2.42 | 2.48 | 2.12 | 2.3 | 2.54 | 2.2 | 2.36 | < 10 | 2.4 | 2.13 | < 10 | 2.5 | 2.34 | < 0.1 | 2.6 | 2.08 | 2.61 | 2.55 | < 10 | 2.09 | 2.43 | 2.4 | [6 empty] | 0
| 1 | 3 | 2.35 | 2.6352 | 2.52 | 2.46 | 2.16 | 2.39 | 2.56 | 2.15 | 2.21 | < 10 | 2.5 | 2.18 | < 10 | 2.53 | 2.48 | < 0.1 | 2.6 | 2.06 | 2.51 | 2.6 | < 10 | 2.16 | 2.43 | 2 | [6 empty] | 16
| 1 | 4 | 2.28 | 2.5759 | 2.46 | 2.44 | 2.2 | 2.35 | 2.57 | 2 | 2.29 | < 10 | 2.5 | 2.16 | < 10 | 2.55 | 2.39 | < 0.1 | 2.7 | 2.06 | 2.52 | 2.66 | < 10 | 2.2 | 2.62 | 2.3 | [6 empty] | 2.3721
| 1 | 5 | 2.25 | 2.654 | 2.56 | 2.54 | 2.08 | 2.28 | 2.57 | 2.2 | 2.37 | < 10 | 2.4 | 2.19 | < 10 | 2.2 | 2.36 | < 0.1 | 2.7 | 2.12 | 2.42 | 2.36 | < 10 | 2.17 | 2.54 | 2 | [6 empty] | 169
| 1 | 6 | 2.35 | 2.6827 | 2.54 | 2.39 | 2.09 | 2.33 | 2.52 | 2.1 | 2.27 | < 10 | 2.5 | 1.91 | < 10 | 2.45 | 2.45 | < 0.1 | 2.7 | 2.04 | 2.49 | 1.81 | < 10 | 2.17 | 2.44 | 2
| Mean |  | 2.285 | 2.6388 | 2.5017 | 2.4517 | 2.1383 | 2.325 | 2.55 | 2.1417 | 2.3017 |  | 2.4667 | 2.1017 |  | 2.4633 | 2.4033 |  | 2.6333 | 2.06 | 2.5233 | 2.425 |  | 2.14 | 2.4933 | 2.1833
| Median |  | 2.295 | 2.6446 | 2.515 | 2.45 | 2.14 | 2.315 | 2.55 | 2.175 | 2.3 |  | 2.5 | 2.145 |  | 2.515 | 2.395 |  | 2.65 | 2.06 | 2.515 | 2.56 |  | 2.165 | 2.47 | 2.15
| Std Dev. |  | 0.068629 | 0.036232 | 0.052313 | 0.055287 | 0.04916 | 0.040373 | 0.02 | 0.080104 | 0.05947 |  | 0.05164 | 0.10834 |  | 0.13441 | 0.053166 |  | 0.08165 | 0.04 | 0.069186 | 0.31779 |  | 0.057271 | 0.076333 | 0.20412
| Rel.Std.Dev. |  | 3.003% | 1.373% | 2.091% | 2.255% | 2.299% | 1.736% | 0.7843% | 3.74% | 2.584% |  | 2.094% | 5.155% |  | 5.457% | 2.212% |  | 3.101% | 1.942% | 2.742% | 13.1% |  | 2.676% | 3.061% | 9.349%
| PDM3 |  | -3.67% | 11.24% | 5.464% | 3.356% | -9.853% | -1.984% | 7.502% | -9.713% | -2.967% |  | 3.989% | -11.4% |  | 3.848% | 1.319% |  | 11.01% | -13.16% | 6.377% | 2.232% |  | -9.783% | 5.113% | -7.956%
| Z-Score (Absolute) |  | 0.77 | 1.15 | 0.4 | 0.13 | 1.57 | 0.56 | 0.67 | 1.55 | 0.68 | NA | 0.21 | 1.56 | NA | 0.48 | 0.13 | NA | 1.12 | 2 | 0.52 | 0.65 | NA | 1.56 | 0.36 | 1.33
| Analytical results for V in OREAS 905 (Certified Value 6.05 ppm) | [31 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 6 | 5.6637 | 7 | 5 | 5 | 6 | 5 | < 10 | 6 | 7 | 6 | 5 | 7 | 6 | 7 | 6 | 7.8 | 7.3 | < 0.1 | 5 | 4 | 5 | 8 | 6 | 5 | 7 | 6 | 7.42 | 5 |  | 1
| 1 | 2 | 6 | 5.5908 | 6 | 6 | 5 | 6 | 5 | < 10 | 6 | 7 | 6 | 5 | 7 | 6 | 7 | 6 | 7.5 | 7.3 | < 0.1 | 6 | 6 | 5 | 6 | 6 | 5 | 7 | 6 | 7.13 | 5 |  | 0
| 1 | 3 | 6 | 5.6238 | 6 | 5 | 5 | 6 | 5 | < 10 | 6 | 7 | 6 | 5 | 7 | 6 | 7 | 6 | 7.3 | 7.1 | < 0.1 | 6 | 6 | 5 | 6 | 6 | 5 | 7 | 6 | 7.23 | 5 |  | 16
| 1 | 4 | 6 | 5.059 | 6 | 5 | 5 | 6 | 5 | < 10 | 6 | 7 | 6 | 5 | 7 | 6 | 7 | 6 | 7.6 | 7.2 | < 0.1 | 6 | 6 | 5 | 6 | 6 | 5 | 7 | 6 | 7.26 | 5 |  | 6.0549
| 1 | 5 | 6 | 5.3428 | 6 | 5 | 5 | 5 | 6 | < 10 | 6 | 7 | 6 | 6 | 7 | 6 | 7 | 6 | 7.7 | 7.8 | < 0.1 | 6 | 6 | 5 | 8 | 6 | 5 | 7 | 6 | 7.36 | 5 |  | 170
| 1 | 6 | 6 | 5.1356 | 6 | 5 | 5 | 6 | 6 | < 10 | 6 | 7 | 6 | 5 | 7 | 6 | 8 | 6 | 7.7 | 7.2 | < 0.1 | 6 | 6 | 5 | 6 | 6 | 5 | 7 | 6 | 7.16 | 5
| Mean |  | 6 | 5.4026 | 6.1667 | 5.1667 | 5 | 5.8333 | 5.3333 |  | 6 | 7 | 6 | 5.1667 | 7 | 6 | 7.1667 | 6 | 7.6 | 7.3167 |  | 5.8333 | 5.6667 | 5 | 6.6667 | 6 | 5 | 7 | 6 | 7.26 | 5
| Median |  | 6 | 5.4668 | 6 | 5 | 5 | 6 | 5 |  | 6 | 7 | 6 | 5 | 7 | 6 | 7 | 6 | 7.65 | 7.25 |  | 6 | 6 | 5 | 6 | 6 | 5 | 7 | 6 | 7.245 | 5
| Std Dev. |  | 0 | 0.26286 | 0.40825 | 0.40825 | 0 | 0.40825 | 0.5164 |  | 0 | 0 | 0 | 0.40825 | 0 | 0 | 0.40825 | 0 | 0.17889 | 0.24833 |  | 0.40825 | 0.8165 | 0 | 1.0328 | 0 | 0 | 0 | 0 | 0.11261 | 0
| Rel.Std.Dev. |  | 0% | 4.865% | 6.62% | 7.902% | 0% | 6.999% | 9.682% |  | 0% | 0% | 0% | 7.902% | 0% | 0% | 5.696% | 0% | 2.354% | 3.394% |  | 6.999% | 14.41% | 0% | 15.49% | 0% | 0% | 0% | 0% | 1.551% | 0%
| PDM3 |  | -0.9069% | -10.77% | 1.846% | -14.67% | -17.42% | -3.659% | -11.92% |  | -0.9069% | 15.61% | -0.9069% | -14.67% | 15.61% | -0.9069% | 18.36% | -0.9069% | 25.52% | 20.84% |  | -3.659% | -6.412% | -17.42% | 10.1% | -0.9069% | -17.42% | 15.61% | -0.9069% | 19.9% | -17.42%
| Z-Score (Absolute) |  | 0 | 0.6 | 0.17 | 0.84 | 1.01 | 0.17 | 0.67 | NA | 0 | 1.01 | 0 | 0.84 | 1.01 | 0 | 1.18 | 0 | 1.62 | 1.23 | NA | 0.17 | 0.34 | 1.01 | 0.67 | 0 | 1.01 | 1.01 | 0 | 1.27 | 1.01
| Analytical results for W in OREAS 905 (Certified Value < 1 ppm) | [31 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 2
| 1 | 1 | 0.61 | 0.52048 | 0.8 | 0.8 | 0.58 | 0.62 | 0.6 | < 0.5 | 0.4 | < 10 | 0.6 | 0.3 | < 10 | 0.5 | < 10 | 0.62 | < 1 | < 5 | 0.5 | < 2 | 0.6 | < 2 | 0.3 | 0.64 | 0.61 | 1.2 | 0.75 | 0.6 | 0.4 |  | 1
| 1 | 2 | 0.58 | 0.46998 | 0.8 | 0.9 | 0.57 | 0.58 | 0.53 | < 0.5 | 0.4 | < 10 | 0.5 | 0.4 | < 10 | 0.5 | < 10 | 0.6 | < 1 | < 5 | 0.4 | < 2 | 0.7 | < 2 | 0.4 | 0.6 | 0.66 | 1.2 | 0.69 | 0.7 | 0.3 |  | 0
| 1 | 3 | 0.6 | 0.50128 | 0.8 | 0.9 | 0.6 | 0.61 | 0.62 | < 0.5 | 0.3 | < 10 | 0.5 | 0.3 | < 10 | 0.5 | < 10 | 0.61 | < 1 | < 5 | 0.5 | < 2 | 0.6 | < 2 | 0.4 | 0.6 | 0.66 | 1.2 | 0.68 | 0.7 | 0.4 |  | 16
| 1 | 4 | 0.63 | 0.45301 | 0.8 | 0.9 | 0.61 | 0.58 | 0.54 | < 0.5 | 0.3 | < 10 | 0.6 | 0.3 | < 10 | 0.5 | < 10 | 0.69 | < 1 | < 5 | 0.5 | < 2 | 0.6 | < 2 | 0.4 | 0.61 | 0.65 | 1.2 | 0.72 | 0.8 | 0.4 |  | < 1
| 1 | 5 | 0.57 | 0.48462 | 0.8 | 1.2 | 0.56 | 0.6 | 0.5 | < 0.5 | 0.3 | < 10 | 0.6 | 0.3 | < 10 | 0.4 | < 10 | 0.58 | < 1 | < 5 | 0.4 | < 2 | 0.7 | < 2 | 0.4 | 0.56 | 0.68 | 1.1 | 0.71 | 0.7 | 0.3 |  | 171
| 1 | 6 | 0.6 | 0.45941 | 0.7 | 0.9 | 0.59 | 0.61 | 0.56 | < 0.5 | 0.2 | < 10 | 0.6 | 0.5 | < 10 | 0.5 | < 10 | 0.63 | < 1 | < 5 | 0.5 | < 2 | 0.6 | < 2 | 0.4 | 0.58 | 0.51 | 1.1 | 0.72 | 0.8 | 0.3
| Mean |  | 0.59833 | 0.48146 | 0.78333 | 0.93333 | 0.585 | 0.6 | 0.55833 |  | 0.31667 |  | 0.56667 | 0.35 |  | 0.48333 |  | 0.62167 |  |  | 0.46667 |  | 0.63333 |  | 0.38333 | 0.59833 | 0.62833 | 1.1667 | 0.71167 | 0.71667 | 0.35
| Median |  | 0.6 | 0.4773 | 0.8 | 0.9 | 0.585 | 0.605 | 0.55 |  | 0.3 |  | 0.6 | 0.3 |  | 0.5 |  | 0.615 |  |  | 0.5 |  | 0.6 |  | 0.4 | 0.6 | 0.655 | 1.2 | 0.715 | 0.7 | 0.35
| Std Dev. |  | 0.02137 | 0.025896 | 0.040825 | 0.13663 | 0.018708 | 0.016733 | 0.044907 |  | 0.075277 |  | 0.05164 | 0.083666 |  | 0.040825 |  | 0.037639 |  |  | 0.05164 |  | 0.05164 |  | 0.040825 | 0.027142 | 0.062423 | 0.05164 | 0.024833 | 0.075277 | 0.054772
| Rel.Std.Dev. |  | 3.572% | 5.379% | 5.212% | 14.64% | 3.198% | 2.789% | 8.043% |  | 23.77% |  | 9.113% | 23.9% |  | 8.447% |  | 6.054% |  |  | 11.07% |  | 8.154% |  | 10.65% | 4.536% | 9.935% | 4.426% | 3.489% | 10.5% | 15.65%
PDM3
| Z-Score (Absolute) |  | 0 | 0.69 | 1.08 | 1.65 | 0.08 | 0.01 | 0.23 | NA | 1.65 | NA | 0.19 | 1.46 | NA | 0.67 | NA | 0.14 | NA | NA | 0.77 | NA | 0.21 | NA | 1.26 | 0 | 0.31 | 3.33 | 0.66 | 0.69 | 1.46
| Analytical results for Y in OREAS 905 (Certified Value 7.08 ppm) | [31 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 18 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [5 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | [5 empty] | 2
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [5 empty] | 3
| 1 | 1 | 7.04 | 6.8173 | 7.31 | 6.8 | 7.17 | 7.68 | 7.32 | 6.75 | 7.39 | 7 | 7 | 6.13 | 7 | 7.39 | 7.34 | 5.5 | 7.3 | 3.8 | 7.4 | 6.41 | 8.28 | 7.43 | 6 | 6.88 | 7.69 | [5 empty] | 1
| 1 | 2 | 6.52 | 6.71 | 6.95 | 6.8 | 7.02 | 7.71 | 7.67 | 6.75 | 7.52 | 7 | 7 | 6.37 | 6 | 7.41 | 7.17 | 5.5 | 7.3 | 4.1 | 7.4 | 6.65 | 7.83 | 7.43 | 6 | 6.61 | 7.56 | [5 empty] | 0
| 1 | 3 | 7.06 | 6.6753 | 7.03 | 6.87 | 7.17 | 7.5 | 7.28 | 6.85 | 7.19 | 7 | 7 | 6.51 | 6 | 7.65 | 7.4 | 5.3 | 7.2 | 4.2 | 7.6 | 3.73 | 7.68 | 7.7 | 6 | 6.91 | 7.55 | [5 empty] | 16
| 1 | 4 | 6.86 | 6.6332 | 7.13 | 6.8 | 7.36 | 7.61 | 7.3 | 6.65 | 7.23 | 7 | 7 | 6.39 | 6 | 7.65 | 7.36 | 5.6 | 7.4 | 4 | 7.7 | 8.75 | 7.45 | 7.92 | 6 | 7.1 | 7.37 | [5 empty] | 7.0814
| 1 | 5 | 6.79 | 6.9491 | 7.03 | 6.9 | 7.13 | 7.4 | 7.37 | 7 | 7.55 | 7 | 7 | 6.71 | 7 | 7.05 | 7.15 | 5.6 | 7.1 | 4.1 | 7.8 | 6.78 | 7.59 | 7.17 | 6 | 6.74 | 7.49 | [5 empty] | 172
| 1 | 6 | 7.04 | 6.8232 | 7.19 | 6.93 | 7.04 | 7.58 | 7.37 | 6.8 | 7.25 | 7 | 7 | 5.92 | 6 | 7.69 | 7.24 | 5.7 | 7.3 | 4.4 | 7.7 | 7.63 | 7.59 | 6.19 | 6 | 6.68 | 7.58
| Mean |  | 6.885 | 6.768 | 7.1067 | 6.85 | 7.1483 | 7.58 | 7.385 | 6.8 | 7.355 | 7 | 7 | 6.3383 | 6.3333 | 7.4733 | 7.2767 | 5.5333 | 7.2667 | 4.1 | 7.6 | 6.6583 | 7.7367 | 7.3067 | 6 | 6.82 | 7.54
| Median |  | 6.95 | 6.7636 | 7.08 | 6.835 | 7.15 | 7.595 | 7.345 | 6.775 | 7.32 | 7 | 7 | 6.38 | 6 | 7.53 | 7.29 | 5.55 | 7.3 | 4.1 | 7.65 | 6.715 | 7.635 | 7.43 | 6 | 6.81 | 7.555
| Std Dev. |  | 0.2105 | 0.11699 | 0.13049 | 0.057966 | 0.12189 | 0.11541 | 0.14433 | 0.11832 | 0.15515 | 0 | 0 | 0.27917 | 0.5164 | 0.24476 | 0.10482 | 0.13663 | 0.10328 | 0.2 | 0.16733 | 1.6716 | 0.29405 | 0.60447 | 0 | 0.179 | 0.10583
| Rel.Std.Dev. |  | 3.057% | 1.729% | 1.836% | 0.8462% | 1.705% | 1.523% | 1.954% | 1.74% | 2.109% | 0% | 0% | 4.404% | 8.154% | 3.275% | 1.44% | 2.469% | 1.421% | 4.878% | 2.202% | 25.11% | 3.801% | 8.273% | 0% | 2.625% | 1.404%
| PDM3 |  | -2.773% | -4.425% | 0.3573% | -3.267% | 0.9457% | 7.042% | 4.288% | -3.973% | 3.864% | -1.149% | -1.149% | -10.49% | -10.56% | 5.535% | 2.758% | -21.86% | 2.617% | -42.1% | 7.324% | -5.974% | 9.254% | 3.182% | -15.27% | -3.691% | 6.477%
| Z-Score (Absolute) |  | 0.49 | 0.74 | 0 | 0.56 | 0.09 | 1.04 | 0.49 | 0.67 | 0.55 | 0.23 | 0.23 | 1.69 | 1.7 | 0.81 | 0.37 | 3.46 | 0.35 | 6.61 | 1.08 | 0.3 | 1.15 | 0.93 | 2.43 | 0.63 | 0.95
| Analytical results for Yb in OREAS 905 (Certified Value 0.27 ppm) | [31 empty] | Certified
Yb | Round | Replicate | 01: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 14: Lab | [5 empty] | 1
| No | No | 01: 03; 03: 04; 04: 05; 05: 09; 06: 10; 07: 12; 08: 13; 09: 15; 10: 17; 11: 22; 12: 26; 13: 30; 14: 32 | [5 empty] | ppm
|  |  | 01: AR*OES/MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*MS; 11: AR*OES/MS; 12: AR*MS; 13: AR*MS; 14: AR*OES/MS | [5 empty] | 2
|  |  | 01: 20g; 03: 0.25g; 04: 0.4g; 05: 40g; 06: 0.25g; 07: 0.5g; 08: 25g; 09: 0.25g; 10: 0.5g; 11: 25g; 12: 40g; 13: 0.4g; 14: 0.25g | [5 empty] | 2
| 1 | 1 | 01: 0.29687; 03: 0.2; 04: 0.3; 05: 0.26; 06: 0.3; 07: 0.3; 08: 0.22; 09: 0.3; 10: 0.27; 11: < 0.1; 12: 0.22; 13: 0.5; 14: 0.3 | [5 empty] | 1
| 1 | 2 | 01: 0.32268; 03: 0.3; 04: 0.3; 05: 0.26; 06: 0.3; 07: 0.2; 08: 0.21; 09: 0.3; 10: 0.27; 11: 0.1; 12: 0.28; 13: 0.5; 14: 0.2 | [5 empty] | 0
| 1 | 3 | 01: 0.28659; 03: 0.3; 04: 0.3; 05: 0.24; 06: 0.3; 07: 0.2; 08: 0.22; 09: 0.3; 10: 0.27; 11: 0.1; 12: 0.25; 13: 0.5; 14: 0.3 | [5 empty] | 16
| 1 | 4 | 01: 0.315; 03: 0.2; 04: 0.3; 05: 0.24; 06: 0.3; 07: 0.3; 08: 0.21; 09: 0.3; 10: 0.24; 11: 0.1; 12: 0.27; 13: 0.5; 14: 0.2 | [5 empty] | 0.26733
| 1 | 5 | 01: 0.30766; 03: 0.2; 04: 0.3; 05: 0.24; 06: 0.3; 07: 0.2; 08: 0.22; 09: 0.3; 10: 0.26; 11: 0.1; 12: 0.22; 13: 0.5; 14: 0.3 | [5 empty] | 173
| 1 | 6 | 01: 0.32526; 03: 0.3; 04: 0.3; 05: 0.24; 06: 0.3; 07: 0.3; 08: 0.2; 09: 0.3; 10: 0.28; 11: 0.1; 12: 0.27; 13: 0.5; 14: 0.2
| Mean |  | 01: 0.30901; 03: 0.25; 04: 0.3; 05: 0.24667; 06: 0.3; 07: 0.25; 08: 0.21333; 09: 0.3; 10: 0.265; 11: 0.1; 12: 0.25167; 13: 0.5; 14: 0.25
| Median |  | 01: 0.31133; 03: 0.25; 04: 0.3; 05: 0.24; 06: 0.3; 07: 0.25; 08: 0.215; 09: 0.3; 10: 0.27; 11: 0.1; 12: 0.26; 13: 0.5; 14: 0.25
| Std Dev. |  | 01: 0.015092; 03: 0.054772; 04: 0; 05: 0.010328; 06: 0; 07: 0.054772; 08: 0.008165; 09: 0; 10: 0.013784; 11: 0; 12: 0.026394; 13: 0; 14: 0.054772
| Rel.Std.Dev. |  | 01: 4.884%; 03: 21.91%; 04: 0%; 05: 4.187%; 06: 0%; 07: 21.91%; 08: 3.827%; 09: 0%; 10: 5.202%; 11: 0%; 12: 10.49%; 13: 0%; 14: 21.91%
| PDM3 |  | 01: 15.59%; 03: -6.484%; 04: 12.22%; 05: -7.731%; 06: 12.22%; 07: -6.484%; 08: -20.2%; 09: 12.22%; 10: -0.8731%; 11: -62.59%; 12: -5.861%; 13: 87.03%; 14: -6.484%
| Z-Score (Absolute) |  | 01: 1.01; 03: 0.03; 04: 0.85; 05: 0.09; 06: 0.85; 07: 0.03; 08: 0.67; 09: 0.85; 10: 0.32; 11: 2.67; 12: 0; 13: 4.37; 14: 0.03
| Analytical results for Zn in OREAS 905 (Certified Value 67 ppm) | [31 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 22 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 |  | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS |  | 0
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.25g | 0.5g | 0.2g | 0.5g | 25g | 0.5g | 0.4g | 0.5g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | 0.5g |  | 1
| 1 | 1 | 69 | 71.795 | 72 | 81 | 65 | 66 | 78 | 76 | 71 | < 50000 | 65.4 | 61 | 67 | 78 | 141 | 63 | 60.4 | 68 | 45 | 60 | 67 | 107 | 61 | 67 | 63 | 63 | 65 | 72.9 | 63.9 |  | 1
| 1 | 2 | 68 | 72.907 | 68 | 80 | 64 | 68 | 80 | 78 | 72 | < 50000 | 62.3 | 62 | 65 | 75 | 145 | 61 | 58.8 | 67 | 46 | 63 | 69 | 76 | 60 | 66 | 62 | 62 | 63 | 65.8 | 66.9 |  | 0
| 1 | 3 | 69 | 70.0783 | 71 | 76 | 65 | 66 | 81 | 77 | 69 | < 50000 | 65.1 | 62 | 67 | 79 | 110 | 65 | 58.7 | 66 | 48 | 63 | 71 | 74 | 59 | 67 | 62 | 63 | 64 | 64.6 | 65.7 |  | 16
| 1 | 4 | 69 | 69.2662 | 70 | 77 | 63 | 67 | 80 | 79 | 74 | < 50000 | 65.6 | 62 | 68 | 73 | 126 | 63 | 61.3 | 68 | 45 | 64 | 71 | 70 | 59 | 65 | 62 | 63 | 66 | 65.5 | 68.8 |  | 67.1665
| 1 | 5 | 70 | 72.7002 | 69 | 75 | 64 | 65 | 82 | 77 | 70 | < 50000 | 65.1 | 66 | 63 | 73 | 114 | 62 | 61.3 | 65 | 46 | 63 | 71 | 68 | 63 | 65 | 62 | 63 | 64 | 64 | 68.2 |  | 174
| 1 | 6 | 70 | 70.923 | 69 | 76 | 63 | 66 | 80 | 78 | 70 | < 50000 | 63.1 | 76 | 66 | 72 | 100 | 64 | 61.8 | 70 | 49 | 62 | 71 | 68 | 64 | 66 | 62 | 62 | 64 | 66.1 | 72.3
| Mean |  | 69.1667 | 71.2783 | 69.8333 | 77.5 | 64 | 66.3333 | 80.1667 | 77.5 | 71 |  | 64.4333 | 64.8333 | 66 | 75 | 122.6667 | 63 | 60.3833 | 67.3333 | 46.5 | 62.5 | 70 | 77.1667 | 61 | 66 | 62.1667 | 62.6667 | 64.3333 | 66.4833 | 67.6333
| Median |  | 69 | 71.359 | 69.5 | 76.5 | 64 | 66 | 80 | 77.5 | 70.5 |  | 65.1 | 62 | 66.5 | 74 | 120 | 63 | 60.85 | 67.5 | 46 | 63 | 71 | 72 | 60.5 | 66 | 62 | 63 | 64 | 65.65 | 67.55
| Std Dev. |  | 0.75277 | 1.453 | 1.472 | 2.429 | 0.89443 | 1.0328 | 1.3292 | 1.0488 | 1.7889 |  | 1.3794 | 5.7417 | 1.7889 | 2.8983 | 17.8624 | 1.4142 | 1.3438 | 1.7512 | 1.6432 | 1.3784 | 1.6733 | 14.9722 | 2.0976 | 0.89443 | 0.40825 | 0.5164 | 1.0328 | 3.2394 | 2.887
| Rel.Std.Dev. |  | 1.088% | 2.039% | 2.108% | 3.134% | 1.398% | 1.557% | 1.658% | 1.353% | 2.52% |  | 2.141% | 8.856% | 2.71% | 3.864% | 14.56% | 2.245% | 2.225% | 2.601% | 3.534% | 2.205% | 2.39% | 19.4% | 3.439% | 1.355% | 0.6567% | 0.824% | 1.605% | 4.872% | 4.269%
| PDM3 |  | 2.978% | 6.122% | 3.971% | 15.38% | -4.714% | -1.24% | 19.36% | 15.38% | 5.708% |  | -4.069% | -3.474% | -1.737% | 11.66% | 82.63% | -6.203% | -10.1% | 0.2484% | -30.77% | -6.948% | 4.219% | 14.89% | -9.181% | -1.737% | -7.444% | -6.699% | -4.218% | -1.017% | 0.6951%
| Z-Score (Absolute) |  | 0.56 | 0.95 | 0.68 | 2.11 | 0.4 | 0.03 | 2.61 | 2.11 | 0.9 | NA | 0.32 | 0.66 | 0.03 | 1.65 | 10.53 | 0.59 | 1.08 | 0.22 | 3.67 | 0.59 | 0.83 | 0.94 | 0.96 | 0.03 | 0.75 | 0.65 | 0.34 | 0.18 | 0.27
| Analytical results for Zr in OREAS 905 (Certified Value 47.5 ppm) | [31 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 14 | 15 | 17 | 19 | 22 | 24 | 26 | 28 | 29 | 30 | 31 | 32 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES/MS | AR*OES/MS | [6 empty] | 1
|  |  | 0.4g | 20g | 0.25g | 0.4g | 0.5g | 0.5g | 20g | 40g | 0.25g | 0.15g | 0.5g | 25g | 10g | 0.25g | 0.5g | 0.5g | 25g | 0.4g | 40g | 0.5g | 0.5g | 0.4g | 0.4g | 0.25g | [6 empty] | 1
| 1 | 1 | 48.5 | 72.7344 | 55.9 | 42.2 | 43 | 47.8 | 48.5 | 49.5 | 48.5 | 47 | 44 | 6.7 | 43 | 55.9 | 48.18 | 43 | 19 | 49.7 | 9 | 50.7 | 39.6 | 41 | 45.3 | 52.3 | [6 empty] | 1
| 1 | 2 | 45.2 | 72.1948 | 53.3 | 44.6 | 41.4 | 47.4 | 46.9 | 50.5 | 53.5 | 48 | 36 | 7.4 | 41 | 55.2 | 46.45 | 42.9 | 21 | 59 | 7 | 45.2 | 40.4 | 41.6 | 44.9 | 50.8 | [6 empty] | 0
| 1 | 3 | 48.9 | 69.2945 | 54.5 | 46.1 | 43.2 | 47.2 | 46.5 | 49 | 51.3 | 48 | 44 | 6.6 | 44 | 54.9 | 45.01 | 43.1 | 21 | 59.8 | 7 | 47.2 | 43.3 | 43.4 | 44.6 | 51 | [6 empty] | 16
| 1 | 4 | 47.2 | 69.3924 | 55.6 | 46.9 | 43.8 | 46.8 | 49 | 51.5 | 50.6 | 49 | 48 | 7.4 | 43 | 55.3 | 44.74 | 42.7 | 21 | 59.2 | 8 | 44.6 | 42.9 | 46.1 | 47.1 | 51.3 | [6 empty] | 47.4967
| 1 | 5 | 46.7 | 69.8771 | 53.6 | 46.8 | 42.5 | 46.2 | 48.4 | 49 | 54 | 49 | 44 | 6 | 45 | 56.9 | 44.74 | 41.9 | 20 | 59.4 | 9 | 45.2 | 38 | 41.3 | 45.8 | 50.4 | [6 empty] | 175
| 1 | 6 | 48.6 | 71.5182 | 55.3 | 43.7 | 42.2 | 47.3 | 49.5 | 50 | 49.9 | 50 | 47 | 7.1 | 44 | 54.3 | 45.53 | 44.1 | 23 | 58.7 | 8 | 46.4 | 33.6 | 41.2 | 44.9 | 49.7
| Mean |  | 47.5167 | 70.8352 | 54.7 | 45.05 | 42.6833 | 47.1167 | 48.1333 | 49.9167 | 51.3 | 48.5 | 43.8333 | 6.8667 | 43.3333 | 55.4167 | 45.775 | 42.95 | 20.8333 | 57.6333 | 8 | 46.55 | 39.6333 | 42.4333 | 45.4333 | 50.9167
| Median |  | 47.85 | 70.6976 | 54.9 | 45.35 | 42.75 | 47.25 | 48.45 | 49.75 | 50.95 | 48.5 | 44 | 6.9 | 43.5 | 55.25 | 45.27 | 42.95 | 21 | 59.1 | 8 | 45.8 | 40 | 41.45 | 45.1 | 50.9
| Std Dev. |  | 1.4275 | 1.503 | 1.0789 | 1.8833 | 0.84004 | 0.55287 | 1.1843 | 0.9704 | 2.1175 | 1.0488 | 4.2151 | 0.54283 | 1.3663 | 0.89536 | 1.3437 | 0.70922 | 1.3292 | 3.9042 | 0.89443 | 2.2412 | 3.5714 | 1.9967 | 0.91579 | 0.87502
| Rel.Std.Dev. |  | 3.004% | 2.122% | 1.972% | 4.181% | 1.968% | 1.173% | 2.461% | 1.944% | 4.128% | 2.162% | 9.616% | 7.905% | 3.153% | 1.616% | 2.935% | 1.651% | 6.38% | 6.774% | 11.18% | 4.815% | 9.011% | 4.705% | 2.016% | 1.719%
| PDM3 |  | 0.04204% | 49.14% | 15.17% | -5.151% | -10.13% | -0.8001% | 1.34% | 5.095% | 8.008% | 2.112% | -7.713% | -85.54% | -8.766% | 16.67% | -3.625% | -9.573% | -56.14% | 21.34% | -83.16% | -1.993% | -16.56% | -10.66% | -4.344% | 7.2%
| Z-Score (Absolute) |  | 0.38 | 5.01 | 1.81 | 0.1 | 0.57 | 0.31 | 0.51 | 0.86 | 1.14 | 0.58 | 0.04 | 7.68 | 0.44 | 1.95 | 0.06 | 0.52 | 4.91 | 2.71 | 7.45 | 0.03 | 1.18 | 0.77 | 0.03 | 1.06
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C85: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C104: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C122: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C140: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C158: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C176: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C194: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C212: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C230: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C248: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C266: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C284: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C302: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C320: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C338: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C356: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C374: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C392: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C410: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C428: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C446: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C464: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C482: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C500: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C518: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C536: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C554: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C572: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C590: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C608: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C626: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C644: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C662: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C680: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C698: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C716: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C734: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C752: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C770: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C788: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C806: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C824: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C842: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C860: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C878: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C896: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C914: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C932: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C950: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C968: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C986: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1004: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1022: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1040: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1058: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1076: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1094: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1112: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1130: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1148: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1166: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1184: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1202: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1220: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1238: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1256: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## H2SO4 Leach
| Analytical results for Cu-Sol in OREAS 905 (Certified Value 1272 ppm) | [28 empty] | Certified
Cu-Sol | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 16 | 17 | 18 | 19 | 23 | 24 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | ppm
|  |  | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*AAS | AL-5S*OES | 0
0
| 1 | 1 | 1220 | 1170 | 1270.796 | 1380 | 1050 | 1230 | 1240 | 1480 | 1240 | 1330 | 1250 | 1330 | 1321 | 1280 | 1278 | 1150 | 1251 | 1300 | 1200 | 1180 | 1400 | 1390 | 1300 | 1200 | 1280 | 1300 | 1270 | 1
| 1 | 2 | 1190 | 1180 | 1261.716 | 1360 | 1080 | 1280 | 1230 | 1500 | 1270 | 1360 | 1240 | 1310 | 1344 | 1170 | 1316 | 1240 | 1291 | 1300 | 1200 | 1190 | 1400 | 1400 | 1310 | 1200 | 1280 | 1300 | 1280 | 0
| 1 | 3 | 1160 | 1180 | 1288.682 | 1370 | 1060 | 1250 | 1250 | 1490 | 1250 | 1360 | 1240 | 1310 | 1314 | 1170 | 1302 | 1210 | 1295 | 1300 | 1200 | 1160 | 1370 | 1390 | 1310 | 1200 | 1290 | 1300 | 1280 | 16
| 1 | 4 | 1180 | 1160 | 1293.838 | 1370 | 1060 | 1250 | 1210 | 1480 | 1280 | 1340 | 1270 | 1290 | 1339 | 1180 | 1276 | 1170 | 1269 | 1300 | 1200 | 1190 | 1390 | 1420 | 1340 | 1200 | 1280 | 1300 | 1280 | 1271.951
| 1 | 5 | 1200 | 1180 | 1282.467 | 1360 | 1020 | 1260 | 1220 | 1480 | 1280 | 1360 | 1290 | 1290 | 1322 | 1170 | 1306 | 1230 | 1279 | 1300 | 1200 | 1180 | 1370 | 1410 | 1320 | 1200 | 1280 | 1300 | 1280 | 177
| 1 | 6 | 1200 | 1180 | 1261.155 | 1350 | 1040 | 1260 | 1240 | 1460 | 1320 | 1350 | 1260 | 1310 | 1307 | 1240 | 1310 | 1170 | 1228 | 1300 | 1300 | 1170 | 1380 | 1380 | 1260 | 1200 | 1280 | 1300 | 1270
| Mean |  | 1191.6667 | 1175 | 1276.4423 | 1365 | 1051.6667 | 1255 | 1231.6667 | 1481.6667 | 1273.3333 | 1350 | 1258.3333 | 1306.6667 | 1324.5 | 1201.6667 | 1298 | 1195 | 1268.8333 | 1300 | 1216.6667 | 1178.3333 | 1385 | 1398.3333 | 1306.6667 | 1200 | 1281.6667 | 1300 | 1276.6667
| Median |  | 1195 | 1180 | 1276.6315 | 1365 | 1055 | 1255 | 1235 | 1480 | 1275 | 1355 | 1255 | 1310 | 1321.5 | 1175 | 1304 | 1190 | 1274 | 1300 | 1200 | 1180 | 1385 | 1395 | 1310 | 1200 | 1280 | 1300 | 1280
| Std Dev. |  | 20.4124 | 8.3666 | 13.9367 | 10.4881 | 20.4124 | 16.4317 | 14.7196 | 13.2916 | 28.0476 | 12.6491 | 19.4079 | 15.0555 | 14.3213 | 47.0815 | 16.9234 | 36.7423 | 25.5532 | 0 | 40.8248 | 11.6905 | 13.784 | 14.7196 | 26.5832 | 0 | 4.0825 | 0 | 5.164
| Rel.Std.Dev. |  | 1.713% | 0.7121% | 1.092% | 0.7684% | 1.941% | 1.309% | 1.195% | 0.8971% | 2.203% | 0.937% | 1.542% | 1.152% | 1.081% | 3.918% | 1.304% | 3.075% | 2.014% | 0% | 3.355% | 0.9921% | 0.9952% | 1.053% | 2.034% | 0% | 0.3185% | 0% | 0.4045%
| PDM3 |  | -6.312% | -7.622% | 0.3531% | 7.315% | -17.32% | -1.333% | -3.167% | 16.49% | 0.1087% | 6.136% | -1.071% | 2.729% | 4.131% | -5.526% | 2.048% | -6.05% | -0.2451% | 2.205% | -4.346% | -7.36% | 8.888% | 9.936% | 2.729% | -5.657% | 0.7638% | 2.205% | 0.3707%
| Z-Score (Absolute) |  | 1.19 | 1.42 | 0 | 1.24 | 3.15 | 0.3 | 0.63 | 2.88 | 0.04 | 1.03 | 0.25 | 0.42 | 0.67 | 1.27 | 0.3 | 1.14 | 0.11 | 0.33 | 0.84 | 1.38 | 1.52 | 1.71 | 0.42 | 1.07 | 0.07 | 0.33 | 0
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Pycnometry
| Analytical results for SG in OREAS 905 (Certified Value 2.75 Unity) | [20 empty] | Certified
SG | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 04 | 06 | 07 | 08 | 09 | 12 | 17 | 19 | 21 | 22 | 25 | 26 | 28 | 29 | 30 | 31 | 32 | 33 | 34 | Unity
|  |  | GASPYC | LIQPYC | LIQPYC | GASPYC | LIQPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | GASPYC | LIQPYC | GASPYC | GASPYC | GASPYC | 2
3
| 1 | 1 | 2.7 | 2.85 | 2.7 | 2.72 | 2.701 | 2.75 | 2.84 | 2.77 | 2.78 | 2.75 | 2.581 | 2.72 | 2.74 | 2.7 | 2.78 | 2.71 | 2.78 | 2.75 | 2.45 | 1
| 1 | 2 | 2.69 | 2.85 | 2.75 | 2.72 | 2.703 | 2.75 | 2.81 | 2.77 | 2.77 | 2.75 | 2.624 | 2.7 | 2.75 | 2.77 | 2.83 | 2.69 | 2.77 | 2.75 | 2.45 | 0
| 1 | 3 | 2.71 | 2.82 | 2.71 | 2.72 | 2.699 | 2.75 | 2.85 | 2.76 | 2.78 | 2.76 | 2.558 | 2.72 | 2.74 | 2.67 | 2.8 | 2.69 | 2.77 | 2.75 | 2.54 | 16
| 1 | 4 | 2.7 | 2.78 | 2.72 | 2.73 | 2.708 | 2.75 | 2.8 | 2.77 | 2.77 | 2.77 | 2.541 | 2.7 | 2.75 | 2.73 | 2.8 | 2.69 | 2.77 | 2.74 | 2.5 | 2.7473
| 1 | 5 | 2.72 | 2.78 | 2.76 | 2.71 | 2.697 | 2.75 | 2.79 | 2.74 | 2.76 | 2.73 | 2.544 | 2.73 | 2.74 | 2.71 | 2.79 | 2.71 | 2.77 | 2.75 | 2.51 | 179
| 1 | 6 | 2.72 | 2.81 | 2.68 | 2.7 | 2.715 | 2.75 | 2.81 | 2.78 | 2.78 | 2.75 | 2.538 | 2.75 | 2.75 | 2.64 | 2.79 | 2.69 | 2.78 | 2.75 | 2.52
| Mean |  | 2.7067 | 2.815 | 2.72 | 2.7167 | 2.7038 | 2.75 | 2.8167 | 2.765 | 2.7733 | 2.7517 | 2.5643 | 2.72 | 2.745 | 2.7033 | 2.7983 | 2.6967 | 2.7733 | 2.7483 | 2.495
| Median |  | 2.705 | 2.815 | 2.715 | 2.72 | 2.702 | 2.75 | 2.81 | 2.77 | 2.775 | 2.75 | 2.551 | 2.72 | 2.745 | 2.705 | 2.795 | 2.69 | 2.77 | 2.75 | 2.505
| Std Dev. |  | 0.012111 | 0.031464 | 0.030332 | 0.010328 | 0.0066458 | 0 | 0.023381 | 0.013784 | 0.008165 | 0.013292 | 0.033255 | 0.018974 | 0.0054772 | 0.045461 | 0.017224 | 0.010328 | 0.005164 | 0.0040825 | 0.037283
| Rel.Std.Dev. |  | 0.4474% | 1.118% | 1.115% | 0.3802% | 0.2458% | 0% | 0.8301% | 0.4985% | 0.2944% | 0.483% | 1.297% | 0.6976% | 0.1995% | 1.682% | 0.6155% | 0.383% | 0.1862% | 0.1485% | 1.494%
| PDM3 |  | -1.478% | 2.465% | -0.9931% | -1.114% | -1.582% | 0.09885% | 2.525% | 0.6448% | 0.9482% | 0.1595% | -6.659% | -0.9931% | -0.08315% | -1.6% | 1.858% | -1.842% | 0.9482% | 0.03818% | -9.183%
| Z-Score (Absolute) |  | 0.91 | 1.67 | 0.59 | 0.67 | 0.98 | 0.12 | 1.71 | 0.48 | 0.67 | 0.16 | 4.3 | 0.59 | 0 | 0.99 | 1.27 | 1.15 | 0.67 | 0.08 | 5.95
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 905 (Indicative Value 14.57 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.52 | 1
| 1 | 2 | 14.61 | 17
| 1 | 3 | 14.51 | 16
| 1 | 4 | 14.56 | 14.5667
| 1 | 5 | 14.59 | 23
| 1 | 6 | 14.61
| Mean |  | 14.5667
| Median |  | 14.575
| Std Dev. |  | 0.044121
| Rel.Std.Dev. |  | 0.3029%
| PDM3 |  | -2.22e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 905 (Indicative Value 0.858 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.84 | 1
| 1 | 2 | 0.85 | 18
| 1 | 3 | 0.94 | 16
| 1 | 4 | 0.84 | 0.85833
| 1 | 5 | 0.84 | 24
| 1 | 6 | 0.84
| Mean |  | 0.85833
| Median |  | 0.84
| Std Dev. |  | 0.040208
| Rel.Std.Dev. |  | 4.684%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 905 (Indicative Value 5.88 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 5.91 | 1
| 1 | 2 | 5.9 | 19
| 1 | 3 | 5.76 | 16
| 1 | 4 | 5.91 | 5.8833
| 1 | 5 | 5.91 | 25
| 1 | 6 | 5.91
| Mean |  | 5.8833
| Median |  | 5.91
| Std Dev. |  | 0.060553
| Rel.Std.Dev. |  | 1.029%
| PDM3 |  | 4.441e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 905 (Indicative Value 3.57 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.57 | 1
| 1 | 2 | 3.58 | 20
| 1 | 3 | 3.54 | 16
| 1 | 4 | 3.57 | 3.5717
| 1 | 5 | 3.58 | 26
| 1 | 6 | 3.59
| Mean |  | 3.5717
| Median |  | 3.575
| Std Dev. |  | 0.017224
| Rel.Std.Dev. |  | 0.4822%
| PDM3 |  | -1.11e-13%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 905 (Indicative Value 0.497 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.5 | 1
| 1 | 2 | 0.5 | 17
| 1 | 3 | 0.48 | 16
| 1 | 4 | 0.5 | 0.49667
| 1 | 5 | 0.5 | 23
| 1 | 6 | 0.5
| Mean |  | 0.49667
| Median |  | 0.5
| Std Dev. |  | 0.008165
| Rel.Std.Dev. |  | 1.644%
| PDM3 |  | -6.661e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 905 (Indicative Value 0.05 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.05 | 1
| 1 | 2 | 0.05 | 18
| 1 | 3 | 0.05 | 16
| 1 | 4 | 0.05 | 0.05
| 1 | 5 | 0.05 | 24
| 1 | 6 | 0.05
| Mean |  | 0.05
| Median |  | 0.05
| Std Dev. |  | 7.6012e-18
| Rel.Std.Dev. |  | 1.52e-14%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 905 (Indicative Value 3.4 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.41 | 1
| 1 | 2 | 3.4 | 19
| 1 | 3 | 3.35 | 16
| 1 | 4 | 3.4 | 3.395
| 1 | 5 | 3.4 | 25
| 1 | 6 | 3.41
| Mean |  | 3.395
| Median |  | 3.4
| Std Dev. |  | 0.022583
| Rel.Std.Dev. |  | 0.6652%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 905 (Indicative Value 0.069 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.067 | 1
| 1 | 2 | 0.069 | 20
| 1 | 3 | 0.071 | 16
| 1 | 4 | 0.068 | 0.0685
| 1 | 5 | 0.068 | 26
| 1 | 6 | 0.068
| Mean |  | 0.0685
| Median |  | 0.068
| Std Dev. |  | 0.0013784
| Rel.Std.Dev. |  | 2.012%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 905 (Indicative Value 68.25 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 68.04 | 1
| 1 | 2 | 68.44 | 17
| 1 | 3 | 68.2 | 16
| 1 | 4 | 68.19 | 68.25
| 1 | 5 | 68.21 | 23
| 1 | 6 | 68.42
| Mean |  | 68.25
| Median |  | 68.205
| Std Dev. |  | 0.15284
| Rel.Std.Dev. |  | 0.2239%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 905 (Indicative Value 0.17 wt.%) |  |  | Indicative
SO3 | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.17 | 1
| 1 | 2 | 0.17 | 18
| 1 | 3 | 0.17 | 16
| 1 | 4 | 0.17 | 0.17
| 1 | 5 | 0.17 | 24
| 1 | 6 | 0.17
| Mean |  | 0.17
| Median |  | 0.17
| Std Dev. |  | 0
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 905 (Indicative Value 0.21 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 26 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.21 | 1
| 1 | 2 | 0.21 | 19
| 1 | 3 | 0.2 | 16
| 1 | 4 | 0.21 | 0.21
| 1 | 5 | 0.22 | 25
| 1 | 6 | 0.21
| Mean |  | 0.21
| Median |  | 0.21
| Std Dev. |  | 0.0063246
| Rel.Std.Dev. |  | 3.012%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Thermograv
| Analytical results for LOI1000 in OREAS 905 (Indicative Value 1.86 wt.%) |  |  |  | Indicative
LOI1000 | Round | Replicate | Lab | Lab | 1
| No | No | 11 | 26 | wt.%
|  |  | at 1000°C | at 1000°C | 2
2
| 1 | 1 | 1.87 | 1.85 | 1
| 1 | 2 | 2.14 | 1.85 | 22
| 1 | 3 | 2.1 | 1.85 | 16
| 1 | 4 | 2.03 | 1.86 | 1.8558
| 1 | 5 | 1.81 | 1.78 | 28
| 1 | 6 | 1.31 | 1.82
| Mean |  | 1.8767 | 1.835
| Median |  | 1.95 | 1.85
| Std Dev. |  | 0.30605 | 0.030166
| Rel.Std.Dev. |  | 16.31% | 1.644%
| PDM3 |  | 1.123% | -1.123%
| Z-Score (Absolute) |  | 0.67 | 0.67
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## IRC
| Analytical results for S in OREAS 905 (Indicative Value 0.053 wt.%) |  |  | Indicative
S | Round | Replicate | Lab | 1
| No | No | 10 | wt.%
|  |  | IRC | 3
3
| 1 | 1 | 0.06 | 1
| 1 | 2 | 0.05 | 24
| 1 | 3 | 0.05 | 16
| 1 | 4 | 0.05 | 0.053333
| 1 | 5 | 0.06 | 30
| 1 | 6 | 0.05
| Mean |  | 0.053333
| Median |  | 0.05
| Std Dev. |  | 0.005164
| Rel.Std.Dev. |  | 9.682%
| PDM3 |  | 6.661e-14%
| Z-Score (Absolute) |  | NA
Note on C17: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.