## Abbreviations
Table 1. Abbreviations used for OREAS 501b
Abbreviation | Explanation
< | ‘less than’
IND | indeterminate
NR | not reported
PDM3 | percent deviation of lab mean from corrected mean of means
Rel.Std.Dev. | one relative standard deviation
Std.Dev. | one standard deviation
Method Codes:
4A*MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-MS finish
4A*OES | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES finish
4A*OES/MS | 4-acid (HF-HNO3-HClO4-HCl) digest with ICP-OES or ICP-MS finish as appropriate
AR*AAS | aqua regia digest with AAS finish
AR*MS | aqua regia digest with ICP-MS finish
AR*OES | aqua regia digest with ICP-OES finish
AR*OES/MS | aqua regia digest with ICP-OES or ICP-MS finish as appropriate
AR*SXAAS | aqua regia digest with solvent extraction AAS finish
FA*AAS | fire assay with AAS finish
FA*OES | fire assay with ICP-OES finish
INAA | INAA using a charge weight as deemed appropriate
Legend:- | (for Method Group Tabulated Results omitted in the determination of Certified Values)
Individual | - Outlying individual values
Batch & Limits | - Outlying laboratory batch data sets and superfluous limit values

## Laboratory List
Table 5. Participating Laboratory List used for OREAS 501b
Lab Slot | Laboratory Name and Location
Accurassay, Thunder Bay, Ontario, Canada
Acme (BV), Santiago, Chile
Acme (BV), Vancouver, BC, Canada
Actlabs, Ancaster, Ontario, Canada
ALS, Brisbane, QLD, Australia
ALS, Johannesburg, South Africa
ALS, La Serena, Coquimbo , Chile
ALS, Lima, Peru
ALS, Perth, WA, Australia
ALS, Vancouver, BC, Canada
Amdel (BV), Adelaide, SA, Australia
CIMM TyS S.A., Antofagasta, Chile
Intertek Genalysis, Johannesburg, South Africa
Intertek Genalysis, Perth, WA, Australia
Intertek Testing Services, Cupang, Muntinlupa, Philippines
Intertek Testing Services, Jakarta, Indonesia
Intertek Testing Services, Shunyi, Beijing, China
Newmont Metallurgical Services, Engelwood, Colorado, USA
OMAC, Loughrea, Galway, Ireland
PT Geoservices Ltd, Cikarang, Jakarta Raya, Indonesia
SGS Canada Inc., Vancouver, BC, Canada
SGS de Mexico, Durango, Mexico
SGS Geosol Laboratorios Ltda, Vespasiano, Minas Gerais, Brazil
SGS Mineral Services, Toronto, Ontario, Canada
SGS South Africa Pty Ltd, Booysens, Gauteng, South Africa
Shiva Analyticals Ltd, Bangalore North, Karnataka, India
Ultra Trace Pty Ltd (BV), Perth, WA, Australia

## Certified Values
Table 2. Certified Values, SD's, 95% Confidence and Tolerance Limits for OREAS 501b
Constituent | Certified | SD | 95% Confidence Limits |  | 95% Tolerance Limits
| Value |  | Low | High | Low | High
Pb Fire Assay
Au, Gold (ppm) | 0.24754 | 0.0095103 | 0.24404 | 0.25104 | 0.23694 | 0.25813
4-Acid Digestion
Ag, Silver (ppm) | 0.77838 | 0.12775 | 0.72513 | 0.83162 | 0.7239 | 0.83285
Al, Aluminium (wt.%) | 7.678 | 0.43067 | 7.4956 | 7.8604 | 7.4705 | 7.8854
As, Arsenic (ppm) | 17.8609 | 2.8538 | 17.0673 | 18.6545 | 14.5184 | 21.2033
Ba, Barium (ppm) | 1009.4709 | 58.4189 | 983.0653 | 1035.8765 | 982.8263 | 1036.1154
Be, Beryllium (ppm) | 2.8567 | 0.29633 | 2.7384 | 2.9751 | 2.7021 | 3.0114
Bi, Bismuth (ppm) | 1.5439 | 0.12525 | 1.4772 | 1.6106 | 1.4514 | 1.6363
Ca, Calcium (wt.%) | 2.7108 | 0.10027 | 2.669 | 2.7526 | 2.6333 | 2.7883
Ce, Cerium (ppm) | 67.629 | 6.1421 | 64.6793 | 70.5787 | 64.3643 | 70.8937
Co, Cobalt (ppm) | 15.8251 | 1.3892 | 15.2717 | 16.3785 | 15.2342 | 16.4159
Cr, Chromium (ppm) | 85.5038 | 8.1804 | 82.0202 | 88.9874 | 80.5659 | 90.4417
Cs, Cesium (ppm) | 12.1535 | 0.731 | 11.7655 | 12.5415 | 11.7176 | 12.5894
Cu, Copper (wt.%) | 0.25967 | 0.010603 | 0.25523 | 0.26411 | 0.25222 | 0.26712
Dy, Dysprosium (ppm) | 4.8596 | 0.21797 | 4.657 | 5.0622 | 4.6744 | 5.0448
Er, Erbium (ppm) | 2.5816 | 0.18631 | 2.385 | 2.7782 | 2.4605 | 2.7027
Eu, Europium (ppm) | 1.3872 | 0.065004 | 1.3274 | 1.447 | 1.2874 | 1.487
Fe, Iron (wt.%) | 4.543 | 0.19154 | 4.4586 | 4.6274 | 4.4298 | 4.6561
Ga, Gallium (ppm) | 19.5236 | 0.78277 | 19.2434 | 19.8037 | 18.5852 | 20.4619
Gd, Gadolinium (ppm) | 5.3977 | 0.12983 | 5.2728 | 5.5225 | 5.2268 | 5.5685
Hf, Hafnium (ppm) | 2.5448 | 0.27921 | 2.3984 | 2.6912 | 2.3596 | 2.73
Ho, Holmium (ppm) | 0.92701 | 0.048386 | 0.881 | 0.97302 | 0.86915 | 0.98487
In, Indium (ppm) | 0.21222 | 0.023295 | 0.20089 | 0.22355 | 0.19923 | 0.2252
K, Potassium (wt.%) | 3.1436 | 0.14233 | 3.0821 | 3.2052 | 3.0386 | 3.2487
La, Lanthanum (ppm) | 32.9652 | 3.4301 | 31.5878 | 34.3426 | 31.0504 | 34.88
Li, Lithium (ppm) | 33.1901 | 3.2223 | 31.7057 | 34.6746 | 31.9672 | 34.4131
Lu, Lutetium (ppm) | 0.37503 | 0.031793 | 0.35191 | 0.39815 | 0.353 | 0.39706
Mg, Magnesium (wt.%) | 1.4821 | 0.057464 | 1.4583 | 1.506 | 1.4482 | 1.516
Mn, Manganese (ppm) | 546.0468 | 24.7369 | 535.6594 | 556.4342 | 532.8742 | 559.2194
Mo, Molybdenum (ppm) | 98.9612 | 7.5361 | 95.7574 | 102.165 | 96.0862 | 101.8361
Na, Sodium (wt.%) | 2.0848 | 0.11098 | 2.0369 | 2.1327 | 2.03 | 2.1396
Nb, Niobium (ppm) | 18.0216 | 1.8149 | 17.1289 | 18.9142 | 17.4021 | 18.641
Nd, Neodymium (ppm) | 30.5668 | 1.2968 | 29.3104 | 31.8231 | 28.7903 | 32.3432
Ni, Nickel (ppm) | 41.4972 | 3.5389 | 40.0915 | 42.9029 | 39.1565 | 43.8379
P, Phosphorus (wt.%) | 0.10246 | 0.0038865 | 0.10073 | 0.1042 | 0.09913 | 0.1058
Pb, Lead (ppm) | 23.0455 | 2.0439 | 22.3072 | 23.7838 | 21.612 | 24.479
Pr, Praseodymium (ppm) | 7.9034 | 0.56651 | 7.2161 | 8.5907 | 7.3914 | 8.4154
Rb, Rubidium (ppm) | 184.2465 | 15.6352 | 175.7454 | 192.7475 | 175.4669 | 193.026
S, Sulphur (wt.%) | 0.35356 | 0.02787 | 0.33995 | 0.36716 | 0.34053 | 0.36658
Sb, Antimony (ppm) | 0.87379 | 0.10185 | 0.81765 | 0.92993 | 0.8016 | 0.94598
Sc, Scandium (ppm) | 12.8978 | 0.83771 | 12.5395 | 13.2562 | 12.364 | 13.4317
Sm, Samarium (ppm) | 5.9725 | 0.31956 | 5.6311 | 6.3138 | 5.622 | 6.3229
Sn, Tin (ppm) | 5.5767 | 0.44481 | 5.3554 | 5.7979 | 5.3882 | 5.7651
Sr, Strontium (ppm) | 337.6362 | 15.9098 | 330.5569 | 344.7155 | 329.119 | 346.1535
Ta, Tantalum (ppm) | 1.3833 | 0.2371 | 1.2549 | 1.5117 | 1.2878 | 1.4788
Tb, Terbium (ppm) | 0.78936 | 0.06279 | 0.73588 | 0.84283 | 0.77327 | 0.80545
Te, Tellurium (ppm) | 0.093285 | 0.014965 | 0.085753 | 0.10082 | IND | IND
Th, Thorium (ppm) | 17.9915 | 0.9774 | 17.7798 | 18.2032 | 16.4385 | 19.5445
Ti, Titanium (wt.%) | 0.46633 | 0.017078 | 0.45968 | 0.47297 | 0.45318 | 0.47947
Tl, Thallium (ppm) | 0.90091 | 0.072824 | 0.86014 | 0.94168 | 0.87768 | 0.92415
Tm, Thulium (ppm) | 0.36278 | 0.030499 | 0.33117 | 0.39439 | 0.3396 | 0.38596
U, Uranium (ppm) | 4.8625 | 0.45712 | 4.6354 | 5.0896 | 4.2528 | 5.4722
V, Vanadium (ppm) | 120.6128 | 5.5975 | 118.2851 | 122.9406 | 117.1428 | 124.0829
Y, Yttrium (ppm) | 24.5665 | 2.2902 | 23.4678 | 25.6652 | 23.8385 | 25.2944
Yb, Ytterbium (ppm) | 2.4926 | 0.073907 | 2.4417 | 2.5435 | 2.4122 | 2.5731
Zn, Zinc (ppm) | 88.8358 | 4.3569 | 87.3016 | 90.3701 | 85.4605 | 92.2111
Zr, Zirconium (ppm) | 77.3386 | 9.25 | 73.1141 | 81.5631 | 73.5193 | 81.1578
Aqua Regia Digestion
Ag, Silver (ppm) | 0.72138 | 0.050786 | 0.69858 | 0.74418 | 0.6946 | 0.74816
Al, Aluminium (wt.%) | 1.9941 | 0.14262 | 1.9272 | 2.0609 | 1.9292 | 2.0589
As, Arsenic (ppm) | 18.5704 | 2.0319 | 18.1777 | 18.963 | 17.1664 | 19.9743
Au, Gold (ppm) | 0.24287 | 0.0094801 | 0.23765 | 0.24809 | 0.23115 | 0.25459
Ba, Barium (ppm) | 368.9409 | 22.7069 | 359.6721 | 378.2098 | 355.7028 | 382.1791
Be, Beryllium (ppm) | 0.48939 | 0.056188 | 0.45735 | 0.52144 | 0.44332 | 0.53547
Bi, Bismuth (ppm) | 1.603 | 0.12064 | 1.5477 | 1.6584 | 1.5139 | 1.6922
Ca, Calcium (wt.%) | 0.96588 | 0.075532 | 0.93008 | 1.0017 | 0.94164 | 0.99011
Ce, Cerium (ppm) | 60.5096 | 2.747 | 59.1868 | 61.8325 | 57.2928 | 63.7265
Co, Cobalt (ppm) | 14.8524 | 1.002 | 14.3773 | 15.3274 | 14.3381 | 15.3666
Cr, Chromium (ppm) | 85.2696 | 4.3986 | 83.5382 | 87.001 | 82.1107 | 88.4284
Cs, Cesium (ppm) | 9.9256 | 0.42573 | 9.6676 | 10.1835 | 9.6496 | 10.2015
Cu, Copper (wt.%) | 0.25812 | 0.007129 | 0.25532 | 0.26093 | 0.25072 | 0.26553
Fe, Iron (wt.%) | 4.078 | 0.25749 | 3.9771 | 4.1789 | 3.9567 | 4.1993
Ga, Gallium (ppm) | 8.6957 | 0.63206 | 8.3634 | 9.0281 | 8.3066 | 9.0849
Ge, Germanium (ppm) | 0.19994 | 0.013824 | 0.18719 | 0.2127 | IND | IND
Hf, Hafnium (ppm) | 0.42758 | 0.059346 | 0.3867 | 0.46846 | 0.39634 | 0.45883
In, Indium (ppm) | 0.19344 | 0.01205 | 0.18639 | 0.2005 | 0.18499 | 0.2019
K, Potassium (wt.%) | 1.0848 | 0.069618 | 1.0551 | 1.1145 | 1.0498 | 1.1198
La, Lanthanum (ppm) | 29.3216 | 1.8973 | 28.4696 | 30.1736 | 28.1273 | 30.5158
Li, Lithium (ppm) | 31.7016 | 2.1689 | 30.7541 | 32.6491 | 30.3421 | 33.0611
Lu, Lutetium (ppm) | 0.21483 | 0.017761 | 0.19804 | 0.23161 | 0.20456 | 0.22509
Mg, Magnesium (wt.%) | 1.2074 | 0.046234 | 1.1881 | 1.2266 | 1.178 | 1.2367
Mn, Manganese (ppm) | 392.9451 | 21.7802 | 384.768 | 401.1222 | 383.3891 | 402.5011
Mo, Molybdenum (ppm) | 96.8982 | 6.4814 | 93.81 | 99.9863 | 94.5571 | 99.2393
Na, Sodium (wt.%) | 0.16711 | 0.029668 | 0.15347 | 0.18075 | 0.15818 | 0.17604
Nd, Neodymium (ppm) | 26.181 | 1.4114 | 24.4808 | 27.8812 | 24.6302 | 27.7318
Ni, Nickel (ppm) | 37.4761 | 2.2936 | 36.5818 | 38.3705 | 35.5404 | 39.4118
P, Phosphorus (wt.%) | 0.099658 | 0.0051651 | 0.097479 | 0.10184 | 0.096279 | 0.10304
Pb, Lead (ppm) | 9.418 | 0.63334 | 9.1545 | 9.6814 | 8.9261 | 9.9098
Pr, Praseodymium (ppm) | 7.173 | 0.31687 | 6.8141 | 7.5319 | 6.8078 | 7.5382
Rb, Rubidium (ppm) | 125.8779 | 5.1125 | 123.2016 | 128.5541 | 121.636 | 130.1198
Re, Rhenium (ppm) | 0.0029048 | 0.00044144 | 0.0026894 | 0.0031201 | IND | IND
S, Sulphur (wt.%) | 0.36421 | 0.027263 | 0.35171 | 0.3767 | 0.34558 | 0.38283
Sb, Antimony (ppm) | 0.4856 | 0.059576 | 0.45173 | 0.51948 | 0.44816 | 0.52304
Sc, Scandium (ppm) | 7.0818 | 0.99207 | 6.644 | 7.5196 | 6.875 | 7.2887
Se, Selenium (ppm) | 2.7835 | 0.4768 | 2.5488 | 3.0182 | 2.6183 | 2.9487
Sn, Tin (ppm) | 4.6506 | 0.33282 | 4.4609 | 4.8403 | 4.4933 | 4.8079
Sr, Strontium (ppm) | 60.1103 | 4.443 | 58.0004 | 62.2202 | 57.5256 | 62.6951
Tb, Terbium (ppm) | 0.55787 | 0.039072 | 0.52619 | 0.58955 | 0.5397 | 0.57604
Te, Tellurium (ppm) | 0.077994 | 0.014683 | 0.070823 | 0.085166 | IND | IND
Th, Thorium (ppm) | 17.9353 | 1.217 | 17.4151 | 18.4555 | 16.5941 | 19.2765
Ti, Titanium (wt.%) | 0.33778 | 0.027548 | 0.32556 | 0.34999 | 0.32579 | 0.34977
Tl, Thallium (ppm) | 0.69367 | 0.04415 | 0.67009 | 0.71724 | 0.67142 | 0.71592
U, Uranium (ppm) | 4.5575 | 0.23504 | 4.4729 | 4.6421 | 4.202 | 4.913
V, Vanadium (ppm) | 109.7369 | 6.0077 | 107.2599 | 112.2139 | 106.8726 | 112.6012
W, Tungsten (ppm) | 2.3445 | 0.43223 | 2.1642 | 2.5248 | 2.0032 | 2.6857
Y, Yttrium (ppm) | 16.5936 | 0.95948 | 16.1216 | 17.0656 | 15.996 | 17.1912
Yb, Ytterbium (ppm) | 1.4956 | 0.049794 | 1.4515 | 1.5397 | 1.4375 | 1.5537
Zn, Zinc (ppm) | 79.8572 | 4.3011 | 78.093 | 81.6213 | 77.4811 | 82.2333
Zr, Zirconium (ppm) | 10.8518 | 1.533 | 10.0884 | 11.6151 | 10.3252 | 11.3784

## Indicative Values
Table 3. Indicative Values for OREAS 501b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Pb Fire Assay
Pd | ppb | 6 | Pt | ppb | 6.6667
4-Acid Digestion
B | ppm | 20 | Hg | ppm | 1.2 | Si | wt.% | 29.5667
Cd | ppm | 0.14556 | Re | ppm | 0.0040926 | W | ppm | 3.5654
Ge | ppm | 0.17958 | Se | ppm | 3.6252
Aqua Regia Digestion
B | ppm | 33.8077 | Gd | ppm | 3.8884 | Pt | ppb | < 5
Cd | ppm | 0.1011 | Hg | ppm | 0.018778 | Si | wt.% | 0.038333
Dy | ppm | 3.2221 | Ho | ppm | 0.61229 | Sm | ppm | 4.71
Er | ppm | 1.6983 | Nb | ppm | 1.3469 | Ta | ppm | 0.69643
Eu | ppm | 0.38242 | Pd | ppb | < 10 | Tm | ppm | 0.22896

## Indicative Values (2)
Table 3. Indicative Values for OREAS 501b
Constituent | Unit | Value | Constituent | Unit | Value | Constituent | Unit | Value
Laser Ablation ICP-MS
Ag | ppm | 0.75 | Ho | ppm | 0.91 | Sn | ppm | 7.4
As | ppm | 19.7 | In | ppm | 0.25 | Sr | ppm | 346.5
Ba | ppm | 1045 | La | ppm | 38.65 | Ta | ppm | 1.455
Be | ppm | 2.6 | Lu | ppm | 0.44 | Tb | ppm | 0.815
Bi | ppm | 1.51 | Mn | wt.% | 0.05785 | Te | ppm | 0.25
Cd | ppm | 0.25 | Mo | ppm | 98.8 | Th | ppm | 19
Ce | ppm | 69.05 | Nb | ppm | 18.05 | Ti | wt.% | 0.5025
Co | ppm | 17.25 | Nd | ppm | 30.7 | Tl | ppm | 0.9
Cr | ppm | 114 | Ni | ppm | 45 | Tm | ppm | 0.375
Cs | ppm | 12 | Pb | ppm | 22 | U | ppm | 5.245
Cu | ppm | 2675 | Pr | ppm | 8.51 | V | ppm | 131.5
Dy | ppm | 4.67 | Rb | ppm | 200.5 | W | ppm | 3.65
Er | ppm | 2.835 | Re | ppm | < 0.01 | Y | ppm | 27.55
Eu | ppm | 1.59 | Sb | ppm | 0.9 | Yb | ppm | 2.665
Ga | ppm | 19.25 | Sc | ppm | 14.7 | Zn | ppm | 80
Gd | ppm | 5.41 | Se | ppm | < 5 | Zr | ppm | 251.5
Hf | ppm | 7.405 | Sm | ppm | 6.44
Borate Fusion XRF
Al2O3 | wt.% | 14.985 | Fe2O3 | wt.% | 6.61 | Pb | ppm | 30
As | ppm | 20 | K2O | wt.% | 3.8 | SiO2 | wt.% | 62.16
Ba | ppm | 1000 | MgO | wt.% | 2.57 | Sn | ppm | 20
CaO | wt.% | 3.91 | MnO | wt.% | 0.08 | SO3 | wt.% | 0.888
Co | ppm | 25 | Na2O | wt.% | 2.83 | TiO2 | wt.% | 0.805
Cr | ppm | 100 | Ni | ppm | 50 | U | ppm | 20
Cu | ppm | 2555 | P2O5 | wt.% | 0.243 | Zn | ppm | 95
Thermogravimetry
LOI1000 | wt.% | 1.11

## Performance Gates
Table 4. Pooled-Lab Performance Gates for OREAS 501b
Constituent | Certified Value | Absolute Standard Deviations |  |  |  |  | Relative Standard Deviations |  |  | 5% window
|  | 1SD | 2SD Low | 2SD High | 3SD Low | 3SD High | 1RSD | 2RSD | 3RSD | Low | High
Pb Fire Assay
Au, ppm | 0.24754 | 0.0095103 | 0.22852 | 0.26656 | 0.21901 | 0.27607 | 3.842% | 7.684% | 11.53% | 0.23516 | 0.25991
4-Acid Digestion
Ag, ppm | 0.77838 | 0.12775 | 0.52287 | 1.0339 | 0.39511 | 1.1616 | 16.41% | 32.83% | 49.24% | 0.73946 | 0.8173
Al, wt.% | 7.678 | 0.43067 | 6.8166 | 8.5393 | 6.3859 | 8.97 | 5.609% | 11.22% | 16.83% | 7.2941 | 8.0619
As, ppm | 17.8609 | 2.8538 | 12.1533 | 23.5684 | 9.2995 | 26.4222 | 15.98% | 31.96% | 47.93% | 16.9678 | 18.7539
Ba, ppm | 1009.4709 | 58.4189 | 892.633 | 1126.3087 | 834.2141 | 1184.7277 | 5.787% | 11.57% | 17.36% | 958.9973 | 1059.9444
Be, ppm | 2.8567 | 0.29633 | 2.2641 | 3.4494 | 1.9677 | 3.7457 | 10.37% | 20.75% | 31.12% | 2.7139 | 2.9996
Bi, ppm | 1.5439 | 0.12525 | 1.2934 | 1.7944 | 1.1681 | 1.9196 | 8.112% | 16.22% | 24.34% | 1.4667 | 1.6211
Ca, wt.% | 2.7108 | 0.10027 | 2.5102 | 2.9113 | 2.41 | 3.0116 | 3.699% | 7.398% | 11.1% | 2.5752 | 2.8463
Ce, ppm | 67.629 | 6.1421 | 55.3448 | 79.9132 | 49.2027 | 86.0553 | 9.082% | 18.16% | 27.25% | 64.2476 | 71.0105
Co, ppm | 15.8251 | 1.3892 | 13.0466 | 18.6035 | 11.6574 | 19.9928 | 8.779% | 17.56% | 26.34% | 15.0338 | 16.6163
Cr, ppm | 85.5038 | 8.1804 | 69.143 | 101.8646 | 60.9626 | 110.0449 | 9.567% | 19.13% | 28.7% | 81.2286 | 89.779
Cs, ppm | 12.1535 | 0.731 | 10.6915 | 13.6155 | 9.9605 | 14.3465 | 6.015% | 12.03% | 18.04% | 11.5458 | 12.7612
Cu, wt.% | 0.25967 | 0.010603 | 0.23846 | 0.28088 | 0.22786 | 0.29148 | 4.083% | 8.167% | 12.25% | 0.24669 | 0.27265
Dy, ppm | 4.8596 | 0.21797 | 4.4236 | 5.2955 | 4.2057 | 5.5135 | 4.485% | 8.971% | 13.46% | 4.6166 | 5.1026
Er, ppm | 2.5816 | 0.18631 | 2.209 | 2.9542 | 2.0227 | 3.1405 | 7.217% | 14.43% | 21.65% | 2.4525 | 2.7107
Eu, ppm | 1.3872 | 0.065004 | 1.2572 | 1.5172 | 1.1922 | 1.5822 | 4.686% | 9.372% | 14.06% | 1.3179 | 1.4566
Fe, wt.% | 4.543 | 0.19154 | 4.1599 | 4.9261 | 3.9684 | 5.1176 | 4.216% | 8.432% | 12.65% | 4.3158 | 4.7701
Ga, ppm | 19.5236 | 0.78277 | 17.958 | 21.0891 | 17.1753 | 21.8719 | 4.009% | 8.019% | 12.03% | 18.5474 | 20.4998
Gd, ppm | 5.3977 | 0.12983 | 5.138 | 5.6573 | 5.0082 | 5.7871 | 2.405% | 4.81% | 7.216% | 5.1278 | 5.6675
Hf, ppm | 2.5448 | 0.27921 | 1.9864 | 3.1032 | 1.7072 | 3.3824 | 10.97% | 21.94% | 32.92% | 2.4176 | 2.6721
Ho, ppm | 0.92701 | 0.048386 | 0.83024 | 1.0238 | 0.78185 | 1.0722 | 5.22% | 10.44% | 15.66% | 0.88066 | 0.97336
In, ppm | 0.21222 | 0.023295 | 0.16563 | 0.25881 | 0.14233 | 0.2821 | 10.98% | 21.95% | 32.93% | 0.20161 | 0.22283
K, wt.% | 3.1436 | 0.14233 | 2.859 | 3.4283 | 2.7166 | 3.5706 | 4.528% | 9.055% | 13.58% | 2.9864 | 3.3008
La, ppm | 32.9652 | 3.4301 | 26.105 | 39.8254 | 22.6748 | 43.2555 | 10.41% | 20.81% | 31.22% | 31.3169 | 34.6135
Li, ppm | 33.1901 | 3.2223 | 26.7455 | 39.6348 | 23.5232 | 42.8571 | 9.709% | 19.42% | 29.13% | 31.5306 | 34.8496
Lu, ppm | 0.37503 | 0.031793 | 0.31145 | 0.43862 | 0.27965 | 0.47041 | 8.477% | 16.95% | 25.43% | 0.35628 | 0.39378
Mg, wt.% | 1.4821 | 0.057464 | 1.3672 | 1.5971 | 1.3097 | 1.6545 | 3.877% | 7.754% | 11.63% | 1.408 | 1.5562
Mn, ppm | 546.0468 | 24.7369 | 496.5729 | 595.5207 | 471.836 | 620.2576 | 4.53% | 9.06% | 13.59% | 518.7445 | 573.3492
Mo, ppm | 98.9612 | 7.5361 | 83.8891 | 114.0333 | 76.353 | 121.5693 | 7.615% | 15.23% | 22.85% | 94.0131 | 103.9092
Na, wt.% | 2.0848 | 0.11098 | 1.8628 | 2.3067 | 1.7519 | 2.4177 | 5.323% | 10.65% | 15.97% | 1.9806 | 2.189
Nb, ppm | 18.0216 | 1.8149 | 14.3917 | 21.6515 | 12.5767 | 23.4664 | 10.07% | 20.14% | 30.21% | 17.1205 | 18.9226
Nd, ppm | 30.5668 | 1.2968 | 27.9732 | 33.1604 | 26.6764 | 34.4572 | 4.243% | 8.485% | 12.73% | 29.0384 | 32.0951
Ni, ppm | 41.4972 | 3.5389 | 34.4194 | 48.575 | 30.8805 | 52.1139 | 8.528% | 17.06% | 25.58% | 39.4223 | 43.5721
P, wt.% | 0.10246 | 0.0038865 | 0.094692 | 0.11024 | 0.090805 | 0.11412 | 3.793% | 7.586% | 11.38% | 0.097342 | 0.10759
Pb, ppm | 23.0455 | 2.0439 | 18.9577 | 27.1333 | 16.9138 | 29.1771 | 8.869% | 17.74% | 26.61% | 21.8932 | 24.1978
Pr, ppm | 7.9034 | 0.56651 | 6.7703 | 9.0364 | 6.2038 | 9.6029 | 7.168% | 14.34% | 21.5% | 7.5082 | 8.2985
Rb, ppm | 184.2465 | 15.6352 | 152.9761 | 215.5169 | 137.3409 | 231.1521 | 8.486% | 16.97% | 25.46% | 175.0341 | 193.4588
S, wt.% | 0.35356 | 0.02787 | 0.29782 | 0.40929 | 0.26995 | 0.43716 | 7.883% | 15.77% | 23.65% | 0.33588 | 0.37123
Sb, ppm | 0.87379 | 0.10185 | 0.67009 | 1.0775 | 0.56824 | 1.1793 | 11.66% | 23.31% | 34.97% | 0.8301 | 0.91748
Sc, ppm | 12.8978 | 0.83771 | 11.2224 | 14.5733 | 10.3847 | 15.411 | 6.495% | 12.99% | 19.48% | 12.2529 | 13.5427
Sm, ppm | 5.9725 | 0.31956 | 5.3333 | 6.6116 | 5.0138 | 6.9311 | 5.351% | 10.7% | 16.05% | 5.6738 | 6.2711
Sn, ppm | 5.5767 | 0.44481 | 4.687 | 6.4663 | 4.2422 | 6.9111 | 7.976% | 15.95% | 23.93% | 5.2978 | 5.8555
Sr, ppm | 337.6362 | 15.9098 | 305.8166 | 369.4558 | 289.9068 | 385.3656 | 4.712% | 9.424% | 14.14% | 320.7544 | 354.518
Ta, ppm | 1.3833 | 0.2371 | 0.90909 | 1.8575 | 0.67199 | 2.0946 | 17.14% | 34.28% | 51.42% | 1.3141 | 1.4525
Tb, ppm | 0.78936 | 0.06279 | 0.66378 | 0.91494 | 0.60099 | 0.97773 | 7.955% | 15.91% | 23.86% | 0.74989 | 0.82883
Te, ppm | 0.093285 | 0.014965 | 0.063354 | 0.12322 | 0.048388 | 0.13818 | 16.04% | 32.09% | 48.13% | 0.08862 | 0.097949
Th, ppm | 17.9915 | 0.9774 | 16.0367 | 19.9463 | 15.0593 | 20.9237 | 5.433% | 10.87% | 16.3% | 17.0919 | 18.8911
Ti, wt.% | 0.46633 | 0.017078 | 0.43217 | 0.50048 | 0.41509 | 0.51756 | 3.662% | 7.324% | 10.99% | 0.44301 | 0.48964
Tl, ppm | 0.90091 | 0.072824 | 0.75526 | 1.0466 | 0.68244 | 1.1194 | 8.083% | 16.17% | 24.25% | 0.85587 | 0.94596
Tm, ppm | 0.36278 | 0.030499 | 0.30178 | 0.42378 | 0.27128 | 0.45428 | 8.407% | 16.81% | 25.22% | 0.34464 | 0.38092
U, ppm | 4.8625 | 0.45712 | 3.9483 | 5.7767 | 3.4911 | 6.2338 | 9.401% | 18.8% | 28.2% | 4.6194 | 5.1056
V, ppm | 120.6128 | 5.5975 | 109.4178 | 131.8079 | 103.8203 | 137.4054 | 4.641% | 9.282% | 13.92% | 114.5822 | 126.6435
Y, ppm | 24.5665 | 2.2902 | 19.9861 | 29.1469 | 17.6959 | 31.4371 | 9.322% | 18.64% | 27.97% | 23.3381 | 25.7948
Yb, ppm | 2.4926 | 0.073907 | 2.3448 | 2.6405 | 2.2709 | 2.7144 | 2.965% | 5.93% | 8.895% | 2.368 | 2.6173
Zn, ppm | 88.8358 | 4.3569 | 80.1221 | 97.5496 | 75.7652 | 101.9065 | 4.904% | 9.809% | 14.71% | 84.394 | 93.2776
Zr, ppm | 77.3386 | 9.25 | 58.8385 | 95.8386 | 49.5884 | 105.0887 | 11.96% | 23.92% | 35.88% | 73.4716 | 81.2055
Aqua Regia Digestion
Ag, ppm | 0.72138 | 0.050786 | 0.61981 | 0.82295 | 0.56902 | 0.87374 | 7.04% | 14.08% | 21.12% | 0.68531 | 0.75745
Al, wt.% | 1.9941 | 0.14262 | 1.7088 | 2.2793 | 1.5662 | 2.4219 | 7.152% | 14.3% | 21.46% | 1.8944 | 2.0938
As, ppm | 18.5704 | 2.0319 | 14.5065 | 22.6342 | 12.4746 | 24.6661 | 10.94% | 21.88% | 32.83% | 17.6418 | 19.4989
Au, ppm | 0.24287 | 0.0094801 | 0.22391 | 0.26183 | 0.21443 | 0.27131 | 3.903% | 7.807% | 11.71% | 0.23073 | 0.25501
Ba, ppm | 368.9409 | 22.7069 | 323.5271 | 414.3548 | 300.8201 | 437.0618 | 6.155% | 12.31% | 18.46% | 350.4939 | 387.388
Be, ppm | 0.48939 | 0.056188 | 0.37702 | 0.60177 | 0.32083 | 0.65796 | 11.48% | 22.96% | 34.44% | 0.46492 | 0.51386
Bi, ppm | 1.603 | 0.12064 | 1.3618 | 1.8443 | 1.2411 | 1.965 | 7.526% | 15.05% | 22.58% | 1.5229 | 1.6832
Ca, wt.% | 0.96588 | 0.075532 | 0.81481 | 1.1169 | 0.73928 | 1.1925 | 7.82% | 15.64% | 23.46% | 0.91758 | 1.0142
Ce, ppm | 60.5096 | 2.747 | 55.0156 | 66.0037 | 52.2686 | 68.7507 | 4.54% | 9.08% | 13.62% | 57.4841 | 63.5351
Co, ppm | 14.8524 | 1.002 | 12.8483 | 16.8564 | 11.8463 | 17.8585 | 6.747% | 13.49% | 20.24% | 14.1097 | 15.595
Cr, ppm | 85.2696 | 4.3986 | 76.4724 | 94.0668 | 72.0738 | 98.4654 | 5.158% | 10.32% | 15.48% | 81.0061 | 89.5331
Cs, ppm | 9.9256 | 0.42573 | 9.0741 | 10.777 | 8.6484 | 11.2028 | 4.289% | 8.578% | 12.87% | 9.4293 | 10.4218
Cu, wt.% | 0.25812 | 0.007129 | 0.24386 | 0.27238 | 0.23673 | 0.27951 | 2.762% | 5.524% | 8.286% | 0.24521 | 0.27103
Fe, wt.% | 4.078 | 0.25749 | 3.563 | 4.593 | 3.3055 | 4.8504 | 6.314% | 12.63% | 18.94% | 3.8741 | 4.2819
Ga, ppm | 8.6957 | 0.63206 | 7.4316 | 9.9598 | 6.7995 | 10.5919 | 7.269% | 14.54% | 21.81% | 8.2609 | 9.1305
Ge, ppm | 0.19994 | 0.013824 | 0.1723 | 0.22759 | 0.15847 | 0.24142 | 6.914% | 13.83% | 20.74% | 0.18995 | 0.20994
Hf, ppm | 0.42758 | 0.059346 | 0.30889 | 0.54627 | 0.24955 | 0.60562 | 13.88% | 27.76% | 41.64% | 0.4062 | 0.44896
In, ppm | 0.19344 | 0.01205 | 0.16934 | 0.21754 | 0.1573 | 0.22959 | 6.229% | 12.46% | 18.69% | 0.18377 | 0.20312
K, wt.% | 1.0848 | 0.069618 | 0.94557 | 1.224 | 0.87595 | 1.2937 | 6.418% | 12.84% | 19.25% | 1.0306 | 1.139
La, ppm | 29.3216 | 1.8973 | 25.5269 | 33.1162 | 23.6296 | 35.0135 | 6.471% | 12.94% | 19.41% | 27.8555 | 30.7876
Li, ppm | 31.7016 | 2.1689 | 27.3639 | 36.0394 | 25.195 | 38.2082 | 6.842% | 13.68% | 20.52% | 30.1165 | 33.2867
Lu, ppm | 0.21483 | 0.017761 | 0.1793 | 0.25035 | 0.16154 | 0.26811 | 8.268% | 16.54% | 24.8% | 0.20408 | 0.22557
Mg, wt.% | 1.2074 | 0.046234 | 1.1149 | 1.2998 | 1.0687 | 1.3461 | 3.829% | 7.659% | 11.49% | 1.147 | 1.2677
Mn, ppm | 392.9451 | 21.7802 | 349.3848 | 436.5054 | 327.6046 | 458.2855 | 5.543% | 11.09% | 16.63% | 373.2978 | 412.5923
Mo, ppm | 96.8982 | 6.4814 | 83.9354 | 109.861 | 77.454 | 116.3424 | 6.689% | 13.38% | 20.07% | 92.0533 | 101.7431
Na, wt.% | 0.16711 | 0.029668 | 0.10777 | 0.22645 | 0.078105 | 0.25612 | 17.75% | 35.51% | 53.26% | 0.15875 | 0.17547
Nd, ppm | 26.181 | 1.4114 | 23.3581 | 29.0038 | 21.9467 | 30.4152 | 5.391% | 10.78% | 16.17% | 24.8719 | 27.49
Ni, ppm | 37.4761 | 2.2936 | 32.8889 | 42.0634 | 30.5953 | 44.357 | 6.12% | 12.24% | 18.36% | 35.6023 | 39.3499
P, wt.% | 0.099658 | 0.0051651 | 0.089328 | 0.10999 | 0.084163 | 0.11515 | 5.183% | 10.37% | 15.55% | 0.094675 | 0.10464
Pb, ppm | 9.418 | 0.63334 | 8.1513 | 10.6846 | 7.5179 | 11.318 | 6.725% | 13.45% | 20.17% | 8.9471 | 9.8889
Pr, ppm | 7.173 | 0.31687 | 6.5392 | 7.8067 | 6.2224 | 8.1236 | 4.418% | 8.835% | 13.25% | 6.8143 | 7.5316
Rb, ppm | 125.8779 | 5.1125 | 115.6529 | 136.1029 | 110.5404 | 141.2153 | 4.061% | 8.123% | 12.18% | 119.584 | 132.1718
Re, ppm | 0.0029048 | 0.00044144 | 0.0020219 | 0.0037876 | 0.0015804 | 0.0042291 | 15.2% | 30.39% | 45.59% | 0.0027595 | 0.00305
S, wt.% | 0.36421 | 0.027263 | 0.30968 | 0.41873 | 0.28242 | 0.446 | 7.486% | 14.97% | 22.46% | 0.346 | 0.38242
Sb, ppm | 0.4856 | 0.059576 | 0.36645 | 0.60476 | 0.30687 | 0.66433 | 12.27% | 24.54% | 36.81% | 0.46132 | 0.50988
Sc, ppm | 7.0818 | 0.99207 | 5.0977 | 9.066 | 4.1056 | 10.0581 | 14.01% | 28.02% | 42.03% | 6.7277 | 7.4359
Se, ppm | 2.7835 | 0.4768 | 1.8299 | 3.7371 | 1.3531 | 4.2139 | 17.13% | 34.26% | 51.39% | 2.6443 | 2.9227
Sn, ppm | 4.6506 | 0.33282 | 3.985 | 5.3162 | 3.6521 | 5.6491 | 7.157% | 14.31% | 21.47% | 4.4181 | 4.8831
Sr, ppm | 60.1103 | 4.443 | 51.2243 | 68.9964 | 46.7813 | 73.4394 | 7.391% | 14.78% | 22.17% | 57.1048 | 63.1158
Tb, ppm | 0.55787 | 0.039072 | 0.47973 | 0.63602 | 0.44066 | 0.67509 | 7.004% | 14.01% | 21.01% | 0.52998 | 0.58576
Te, ppm | 0.077994 | 0.014683 | 0.048628 | 0.10736 | 0.033944 | 0.12204 | 18.83% | 37.65% | 56.48% | 0.074095 | 0.081894
Th, ppm | 17.9353 | 1.217 | 15.5014 | 20.3692 | 14.2844 | 21.5862 | 6.785% | 13.57% | 20.36% | 17.0385 | 18.8321
Ti, wt.% | 0.33778 | 0.027548 | 0.28268 | 0.39287 | 0.25513 | 0.42042 | 8.156% | 16.31% | 24.47% | 0.32089 | 0.35467
Tl, ppm | 0.69367 | 0.04415 | 0.60537 | 0.78197 | 0.56122 | 0.82612 | 6.365% | 12.73% | 19.09% | 0.65899 | 0.72835
U, ppm | 4.5575 | 0.23504 | 4.0874 | 5.0276 | 3.8524 | 5.2626 | 5.157% | 10.31% | 15.47% | 4.3296 | 4.7854
V, ppm | 109.7369 | 6.0077 | 97.7214 | 121.7524 | 91.7136 | 127.7601 | 5.475% | 10.95% | 16.42% | 104.25 | 115.2237
W, ppm | 2.3445 | 0.43223 | 1.48 | 3.2089 | 1.0478 | 3.6412 | 18.44% | 36.87% | 55.31% | 2.2273 | 2.4617
Y, ppm | 16.5936 | 0.95948 | 14.6746 | 18.5125 | 13.7151 | 19.472 | 5.782% | 11.56% | 17.35% | 15.7639 | 17.4233
Yb, ppm | 1.4956 | 0.049794 | 1.396 | 1.5952 | 1.3462 | 1.645 | 3.329% | 6.659% | 9.988% | 1.4208 | 1.5704
Zn, ppm | 79.8572 | 4.3011 | 71.255 | 88.4594 | 66.9538 | 92.7606 | 5.386% | 10.77% | 16.16% | 75.8643 | 83.8501
Zr, ppm | 10.8518 | 1.533 | 7.7857 | 13.9179 | 6.2527 | 15.4509 | 14.13% | 28.25% | 42.38% | 10.3092 | 11.3944

## Fire Assay
| Analytical results for Au in OREAS 501b (Certified Value 0.248 ppm) | [27 empty] | Certified
Au | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 00 | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 19 | 20 | 21 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | INAA | FA*AAS | FA*AAS | FA*OES | FA*OES | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*OES | FA*AAS | FA*OES | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | FA*AAS | 3
|  |  | 1.0g | 40g | 30g | 40g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 25g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 30g | 40g | 3
| 1 | 1 | 0.284 | 0.2463 | 0.27 | 0.24 | 0.244 | 0.242 | 0.262 | 0.261 | 0.24 | 0.249 | 0.25 | 0.246 | 0.251 | 0.26 | 0.248 | 0.252 | 0.242 | 0.227 | 0.238 |  | 0.253 | 0.22 | 0.24 | 0.262 | 0.22 | 0.24 | 1
| 1 | 2 | 0.271 | 0.2514 | 0.27 | 0.242 | 0.25 | 0.242 | 0.247 | 0.257 | 0.24 | 0.249 | 0.27 | 0.247 | 0.25 | 0.25 | 0.247 | 0.25026 | 0.24 | 0.172 | 0.247 | 0.24655 | 0.254 | 0.23 | 0.23 | 0.255 | 0.25 | 0.24 | #N/A
| 1 | 3 | 0.266 | 0.2558 | 0.27 | 0.24 | 0.25 | 0.241 | 0.261 | 0.258 | 0.24 | 0.248 | 0.25 | 0.241 | 0.253 | 0.25 | 0.244 | 0.25485 | 0.242 | 0.154 | 0.228 | 0.24673 | 0.258 | 0.24 | 0.25 | 0.26 | 0.24 | 0.24 | 16
| 1 | 4 | 0.257 | 0.2504 | 0.28 | 0.241 | 0.249 | 0.243 | 0.26 | 0.258 | 0.24 | 0.249 | 0.26 | 0.231 | 0.243 | 0.25 | 0.242 | 0.25394 | 0.242 | 0.231 | 0.251 | 0.24765 | 0.259 | 0.23 | 0.24 | 0.261 | 0.25 | 0.25 | 0.24754
| 1 | 5 | 0.252 | 0.2571 | 0.27 | 0.243 | 0.247 | 0.244 | 0.256 | 0.268 | 0.24 | 0.259 | 0.27 | 0.241 | 0.247 | 0.26 | 0.243 | 0.2529 | 0.241 | 0.239 | 0.237 | 0.23753 | 0.257 | 0.24 | 0.24 | 0.261 | 0.23 | 0.24 | 7
| 1 | 6 | 0.258 | 0.248 | 0.27 | 0.243 | 0.248 | 0.235 | 0.26 | 0.253 | 0.27 | 0.26 | 0.27 | 0.263 | 0.247 | 0.25 | 0.242 | 0.24827 | 0.242 | 0.238 | 0.231 | 0.24339 | 0.257 | 0.24 | 0.25 | 0.259 | 0.2 | 0.24
|  | 7 | 00: 0.242
|  | 8 | 00: 0.234
|  | 9 | 00: 0.239
|  | 10 | 00: 0.238
|  | 11 | 00: 0.214
|  | 12 | 00: 0.228
|  | 13 | 00: 0.255
|  | 14 | 00: 0.237
|  | 15 | 00: 0.251
|  | 16 | 00: 0.263
|  | 17 | 00: 0.221
|  | 18 | 00: 0.248
|  | 19 | 00: 0.249
|  | 20 | 00: 0.277
| Mean |  | 0.2492 | 0.2515 | 0.27167 | 0.2415 | 0.248 | 0.24117 | 0.25767 | 0.25917 | 0.245 | 0.25233 | 0.26167 | 0.24483 | 0.2485 | 0.25333 | 0.24433 | 0.25204 | 0.2415 | 0.21017 | 0.23867 | 0.24437 | 0.25633 | 0.23333 | 0.24167 | 0.25967 | 0.23167 | 0.24167
| Median |  | 0.25 | 0.2509 | 0.27 | 0.2415 | 0.2485 | 0.242 | 0.26 | 0.258 | 0.24 | 0.249 | 0.265 | 0.2435 | 0.2485 | 0.25 | 0.2435 | 0.25245 | 0.242 | 0.229 | 0.2375 | 0.24655 | 0.257 | 0.235 | 0.24 | 0.2605 | 0.235 | 0.24
| Std Dev. |  | 0.018153 | 0.0042511 | 0.0040825 | 0.0013784 | 0.0022804 | 0.0031885 | 0.0056095 | 0.0050365 | 0.012247 | 0.0055737 | 0.0098319 | 0.010553 | 0.0035637 | 0.005164 | 0.002582 | 0.002436 | 0.00083666 | 0.037242 | 0.0089144 | 0.0041482 | 0.0023381 | 0.008165 | 0.0075277 | 0.0025033 | 0.019408 | 0.0040825
| Rel.Std.Dev. |  | 7.285% | 1.69% | 1.503% | 0.5708% | 0.9195% | 1.322% | 2.177% | 1.943% | 4.999% | 2.209% | 3.757% | 4.31% | 1.434% | 2.038% | 1.057% | 0.9665% | 0.3464% | 17.72% | 3.735% | 1.698% | 0.9121% | 3.499% | 3.115% | 0.9641% | 8.378% | 1.689%
| PDM3 |  | 0.6715% | 1.601% | 9.748% | -2.439% | 0.1867% | -2.574% | 4.092% | 4.698% | -1.025% | 1.937% | 5.708% | -1.093% | 0.3887% | 2.341% | -1.295% | 1.817% | -2.439% | -15.1% | -3.584% | -1.28% | 3.553% | -5.738% | -2.372% | 4.9% | -6.412% | -2.372%
| Z-Score (Absolute) |  | NA | 0.67 | 2.77 | 0.36 | 0.31 | 0.4 | 1.31 | 1.47 | 0 | 0.76 | 1.73 | 0.02 | 0.36 | 0.86 | 0.07 | 0.73 | 0.36 | 3.61 | 0.66 | 0.07 | 1.18 | 1.21 | 0.35 | 1.52 | 1.38 | 0.35
| Analytical results for Pd in OREAS 501b (Indicative Value 6 ppb) | [27 empty] | Indicative
Pd | Round | Replicate | 00: Lab | 1
| No | No | 00: 03 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 40g | 2
| 1 | 1 | 00: 5 | 1
| 1 | 2 | 00: 5 | 1
| 1 | 3 | 00: 10 | 16
| 1 | 4 | 00: < 5 | 6
| 1 | 5 | 00: 15 | 7
| 1 | 6 | 00: 5
| Mean |  | 00: 8
| Median |  | 00: 5
| Std Dev. |  | 00: 4.4721
| Rel.Std.Dev. |  | 00: 55.9%
| PDM3 |  | 00: 33.33%
Z-Score (Absolute) | bar: 00
| Analytical results for Pt in OREAS 501b (Indicative Value 6.67 ppb) | [27 empty] | Indicative
Pt | Round | Replicate | 00: Lab | 1
| No | No | 00: 03 | ppb
|  |  | 00: FA*OES | 2
|  |  | 00: 40g | 2
| 1 | 1 | 00: 5 | 1
| 1 | 2 | 00: 10 | 1
| 1 | 3 | 00: 5 | 16
| 1 | 4 | 00: 10 | 6.6667
| 1 | 5 | 00: 5 | 7
| 1 | 6 | 00: 5
| Mean |  | 00: 6.6667
| Median |  | 00: 5
| Std Dev. |  | 00: 2.582
| Rel.Std.Dev. |  | 00: 38.73%
| PDM3 |  | 00: -4.441e-14%
Z-Score (Absolute) | bar: 00
Note on C31: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C49: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C67: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Aqua Regia
| Analytical results for Ag in OREAS 501b (Certified Value 0.721 ppm) | [24 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES | 3
3
| 1 | 1 | 0.7253 | 0.65 | 0.65 | 0.72 | 0.76 | 0.8 | 0.64 | 0.75 | 5 | 0.4 | 0.5 | 0.73 | 0.7 | 0.679 | 0.77 | 0.79 | < 1 | 0.74 | 0.64 | 0.7 | 0.5 | 1.1 | 0.79 | 1
| 1 | 2 | 0.7542 | 0.7 | 0.65 | 0.72 | 0.76 | 0.8 | 0.68 | 0.76 | 5 | 0.4 | 0.5 | 0.67 | 0.7 | 0.68 | 0.76 | 0.79 | < 1 | 0.71 | 0.63 | 0.8 | 0.49 | 1 | 0.68 | #N/A
| 1 | 3 | 0.7471 | 0.73 | 0.7 | 0.72 | 0.76 | 0.8 | 0.66 | 0.73 | 6 | 0.4 | 0.5 | 0.68 | 0.7 | 0.695 | 0.77 | 0.76 | < 1 | 0.74 | 0.63 | 0.8 | 0.48 | 0.8 | 0.73 | 16
| 1 | 4 | 0.7115 | 0.63 | 0.7 | 0.72 | 0.78 | 0.8 | 0.67 | 0.73 | 5 | 0.5 | 0.5 | 0.69 | 0.7 | 0.696 | 0.77 | 0.75 | 2 | 0.74 | 0.64 | 0.8 | 0.45 | 1 | 0.84 | 0.72138
| 1 | 5 | 0.7209 | 0.62 | 0.7 | 0.71 | 0.76 | 0.8 | 0.67 | 0.78 | 5 | 0.5 | 0.5 | 0.71 | 0.6 | 0.724 | 0.74 | 0.75 | 1 | 0.74 | 0.66 | 0.7 | 0.49 | 0.9 | 0.66 | 65
| 1 | 6 | 0.7568 | 0.7 | 0.65 | 0.72 | 0.76 | 0.8 | 0.64 | 0.74 | 5 | 0.4 | 0.5 | 0.71 | 0.7 | 0.669 | 0.77 | 0.71 | 1 | 0.79 | 0.65 | 0.8 | 0.41 | 1.1 | 0.7
| Mean |  | 0.73597 | 0.67167 | 0.675 | 0.71833 | 0.76333 | 0.8 | 0.66 | 0.74833 | 5.1667 | 0.43333 | 0.5 | 0.69833 | 0.68333 | 0.6905 | 0.76333 | 0.75833 | 1.3333 | 0.74333 | 0.64167 | 0.76667 | 0.47 | 0.98333 | 0.73333
| Median |  | 0.7362 | 0.675 | 0.675 | 0.72 | 0.76 | 0.8 | 0.665 | 0.745 | 5 | 0.4 | 0.5 | 0.7 | 0.7 | 0.6875 | 0.77 | 0.755 | 1 | 0.74 | 0.64 | 0.8 | 0.485 | 1 | 0.715
| Std Dev. |  | 0.01913 | 0.04446 | 0.027386 | 0.0040825 | 0.008165 | 1.2162e-16 | 0.016733 | 0.019408 | 0.40825 | 0.05164 | 0 | 0.022286 | 0.040825 | 0.019378 | 0.012111 | 0.029944 | 0.57735 | 0.02582 | 0.01169 | 0.05164 | 0.034059 | 0.1169 | 0.069186
| Rel.Std.Dev. |  | 2.599% | 6.619% | 4.057% | 0.5683% | 1.07% | 1.52e-14% | 2.535% | 2.593% | 7.902% | 11.92% | 0% | 3.191% | 5.974% | 2.806% | 1.587% | 3.949% | 43.3% | 3.474% | 1.822% | 6.736% | 7.247% | 11.89% | 9.434%
| PDM3 |  | 2.022% | -6.891% | -6.429% | -0.4224% | 5.816% | 10.9% | -8.509% | 3.736% | 616.2% | -39.93% | -30.69% | -3.195% | -5.274% | -4.281% | 5.816% | 5.123% | 84.83% | 3.043% | -11.05% | 6.278% | -34.85% | 36.31% | 1.657%
| Z-Score (Absolute) |  | 0.04 | 0.83 | 0.79 | 0.2 | 0.4 | 0.9 | 0.99 | 0.2 | 59.79 | 4.05 | 3.15 | 0.47 | 0.67 | 0.58 | 0.4 | 0.34 | 5.84 | 0.13 | 1.24 | 0.45 | 3.55 | 3.37 | 0
| Analytical results for Al in OREAS 501b (Certified Value 1.99 wt.%) | [24 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 2
3
| 1 | 1 | 2.3016 | 2.172 | 1.91 | 1.81 | 1.98 | 1.86 | 1.89 | 1.9192 | 1.36 | 1.91 | 1.88 | 2.14 | 1.96 | 2.0529 | 1.91 | 2.08 | 1.42 | 2.08 | 1.92 | 1.89 | 2.19 | 3.53 | 2.25 | 1
| 1 | 2 | 2.3105 | 2.03 | 1.91 | 1.86 | 1.95 | 1.91 | 1.92 | 1.9481 | 1.34 | 1.94 | 1.93 | 2.11 | 1.91 | 2.0888 | 1.95 | 2.02 | 1.46 | 2.2 | 2.01 | 1.88 | 2.12 | 3.97 | 2.24 | #N/A
| 1 | 3 | 2.2329 | 2.204 | 1.96 | 1.83 | 1.99 | 1.85 | 1.9 | 1.8765 | 1.37 | 1.98 | 1.92 | 2.14 | 1.91 | 2.045 | 1.99 | 2.04 | 1.61 | 2.12 | 1.96 | 1.93 | 2.13 | 3.55 | 2.26 | 16
| 1 | 4 | 2.2424 | 2.193 | 1.95 | 1.84 | 1.96 | 1.86 | 1.93 | 1.8993 | 1.35 | 1.94 | 1.94 | 2.17 | 1.85 | 2.1083 | 1.97 | 2.09 | 1.9 | 2.1 | 1.91 | 1.97 | 2.1 | 3.8 | 2.28 | 1.9941
| 1 | 5 | 2.1817 | 2.179 | 1.88 | 1.88 | 1.97 | 1.83 | 1.89 | 1.9129 | 1.37 | 1.97 | 1.86 | 2.18 | 1.86 | 2.0949 | 1.99 | 2.02 | 1.64 | 2.11 | 1.88 | 1.91 | 2.16 | 2.99 | 2.27 | 66
| 1 | 6 | 2.2514 | 2.151 | 1.9 | 1.88 | 1.95 | 1.86 | 1.92 | 1.8798 | 1.33 | 1.95 | 1.93 | 2.18 | 1.85 | 2.0382 | 2.03 | 2.02 | 1.52 | 2.11 | 1.98 | 2.01 | 2.1 | 4.02 | 2.34
| Mean |  | 2.2534 | 2.1548 | 1.9183 | 1.85 | 1.9667 | 1.8617 | 1.9083 | 1.906 | 1.3533 | 1.9483 | 1.91 | 2.1533 | 1.89 | 2.0714 | 1.9733 | 2.045 | 1.5917 | 2.12 | 1.9433 | 1.9317 | 2.1333 | 3.6433 | 2.2733
| Median |  | 2.2469 | 2.1755 | 1.91 | 1.85 | 1.965 | 1.86 | 1.91 | 1.9061 | 1.355 | 1.945 | 1.925 | 2.155 | 1.885 | 2.0709 | 1.98 | 2.03 | 1.565 | 2.11 | 1.94 | 1.92 | 2.125 | 3.675 | 2.265
| Std Dev. |  | 0.047476 | 0.063798 | 0.030605 | 0.028284 | 0.01633 | 0.026394 | 0.017224 | 0.026817 | 0.01633 | 0.024833 | 0.032249 | 0.028048 | 0.044272 | 0.02951 | 0.040825 | 0.032094 | 0.17302 | 0.041473 | 0.048442 | 0.049967 | 0.03559 | 0.37988 | 0.03559
| Rel.Std.Dev. |  | 2.107% | 2.961% | 1.595% | 1.529% | 0.8303% | 1.418% | 0.9026% | 1.407% | 1.207% | 1.275% | 1.688% | 1.303% | 2.342% | 1.425% | 2.069% | 1.569% | 10.87% | 1.956% | 2.493% | 2.587% | 1.668% | 10.43% | 1.566%
| PDM3 |  | 13.01% | 8.062% | -3.798% | -7.225% | -1.374% | -6.64% | -4.299% | -4.418% | -32.13% | -2.293% | -4.216% | 7.987% | -5.219% | 3.877% | -1.04% | 2.554% | -20.18% | 6.315% | -2.544% | -3.129% | 6.984% | 82.71% | 14%
| Z-Score (Absolute) |  | 2.13 | 1.44 | 0.21 | 0.69 | 0.13 | 0.6 | 0.28 | 0.3 | 4.15 | 0 | 0.27 | 1.43 | 0.41 | 0.86 | 0.17 | 0.67 | 2.49 | 1.2 | 0.03 | 0.12 | 1.29 | 11.82 | 2.27
| Analytical results for As in OREAS 501b (Certified Value 18.6 ppm) | [24 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 1
1
| 1 | 1 | 18.463 | 8 | 18 | 19.4 | 21.9 | 17.9 | 17 | 17 | 13 | 15 | 20 | 15.8 | 16 | 20.3259 | 20.1 | 16.6 | 22 | 23.2 | 20 | 23 | 15 | 9 | 22 | 1
| 1 | 2 | 19.323 | 9 | 18.2 | 19.1 | 29.4 | 22.3 | 19 | 17 | 19 | 15 | 24 | 17.5 | 19 | 14.5368 | 19.6 | 22.5 | 17 | 19 | 17 | 22 | 20 | 7 | 19 | #N/A
| 1 | 3 | 18.986 | 11 | 17.8 | 14.7 | 17.9 | 18.2 | 19 | 20 | 15 | 15 | 20 | 20.2 | 21 | 17.1426 | 15.6 | 18.7 | 19 | 24.6 | 7 | 19 | 16 | 7 | 20 | 16
| 1 | 4 | 17.473 | 8 | 19.2 | 22.4 | 33.8 | 18 | 21 | 19 | 14 | 17 | 20 | 17.6 | 20 | 19.8257 | 15.9 | 21.2 | 22 | 18.9 | 11 | 19 | 18 | 10 | 15 | 18.5704
| 1 | 5 | 18.372 | 11 | 17.8 | 16.4 | 19.8 | 18.8 | 17 | 17 | 17 | 16 | 17 | 20.9 | 18 | 18.2372 | 18.3 | 20 | 17 | 19.1 | 7 | 20 | 17 | 11 | 14 | 67
| 1 | 6 | 19.274 | 10 | 18.2 | 21.5 | 19.8 | 18.5 | 16 | 19 | 16 | 17 | 20 | 18.7 | 18 | 16.5154 | 16.3 | 18.2 | 15 | 19.9 | 9 | 20 | 24 | 15 | 18
| Mean |  | 18.6485 | 9.5 | 18.2 | 18.9167 | 23.7667 | 18.95 | 18.1667 | 18.1667 | 15.6667 | 15.8333 | 20.1667 | 18.45 | 18.6667 | 17.7639 | 17.6333 | 19.5333 | 18.6667 | 20.7833 | 11.8333 | 20.5 | 18.3333 | 9.8333 | 18
| Median |  | 18.7245 | 9.5 | 18.1 | 19.25 | 20.85 | 18.35 | 18 | 18 | 15.5 | 15.5 | 20 | 18.15 | 18.5 | 17.6899 | 17.3 | 19.35 | 18 | 19.5 | 10 | 20 | 17.5 | 9.5 | 18.5
| Std Dev. |  | 0.7004 | 1.3784 | 0.52154 | 2.939 | 6.3525 | 1.6742 | 1.8348 | 1.3292 | 2.1602 | 0.98319 | 2.2286 | 1.8855 | 1.7512 | 2.1632 | 1.9654 | 2.1388 | 2.8752 | 2.4799 | 5.4559 | 1.6432 | 3.266 | 2.9944 | 3.0332
| Rel.Std.Dev. |  | 3.756% | 14.51% | 2.866% | 15.54% | 26.73% | 8.835% | 10.1% | 7.316% | 13.79% | 6.21% | 11.05% | 10.22% | 9.381% | 12.18% | 11.15% | 10.95% | 15.4% | 11.93% | 46.11% | 8.015% | 17.81% | 30.45% | 16.85%
| PDM3 |  | 0.4208% | -48.84% | -1.994% | 1.865% | 27.98% | 2.044% | -2.174% | -2.174% | -15.64% | -14.74% | 8.596% | -0.6481% | 0.5186% | -4.343% | -5.046% | 5.186% | 0.5186% | 11.92% | -36.28% | 10.39% | -1.276% | -47.05% | -3.071%
| Z-Score (Absolute) |  | 0.34 | 9.66 | 0.15 | 0.64 | 5.94 | 0.67 | 0.18 | 0.18 | 2.92 | 2.73 | 2 | 0.13 | 0.36 | 0.62 | 0.77 | 1.31 | 0.36 | 2.68 | 7.11 | 2.37 | 0 | 9.29 | 0.36
| Analytical results for Au in OREAS 501b (Certified Value 0.243 ppm) | [24 empty] | Certified
Au | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 05; 06: 06; 07: 08; 08: 10; 09: 17; 10: 18; 11: 23; 12: 27 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*SXAAS; 08: AR*AAS; 09: AR*MS; 10: AR*MS; 11: AR*AAS; 12: AR*AAS | 3
|  |  | 07: 25g | 3
| 1 | 1 | 01: 0.23747; 02: 0.222; 03: 0.25; 04: 0.248; 05: 0.249; 06: 0.2411; 07: 0.24; 08: 0.24; 09: 0.236; 10: 0.261; 11: 0.24; 12: 0.22 | 1
| 1 | 2 | 01: 0.24479; 02: 0.22; 03: 0.248; 04: 0.237; 05: 0.26; 06: 0.2619; 07: 0.24; 08: 0.24; 09: 0.239; 10: 0.269; 11: 0.26; 12: 0.2 | #N/A
| 1 | 3 | 01: 0.24798; 02: 0.221; 03: 0.254; 04: 0.235; 05: 0.249; 06: 0.2563; 07: 0.23; 08: 0.24; 09: 0.239; 10: 0.261; 11: 0.24; 12: 0.21 | 16
| 1 | 4 | 01: 0.22817; 02: 0.233; 03: 0.243; 04: 0.244; 05: 0.261; 06: 0.2502; 07: 0.24; 08: 0.23; 09: 0.243; 10: 0.267; 11: 0.23; 12: 0.2 | 0.24287
| 1 | 5 | 01: 0.23395; 02: 0.251; 03: 0.251; 04: 0.24; 05: 0.257; 06: 0.2552; 07: 0.24; 08: 0.24; 09: 0.253; 10: 0.251; 11: 0.24; 12: 0.2 | 68
| 1 | 6 | 01: 0.24097; 02: 0.24; 03: 0.249; 04: 0.244; 05: 0.254; 06: 0.2482; 07: 0.24; 08: 0.24; 09: 0.256; 10: 0.263; 11: 0.25; 12: 0.21
| Mean |  | 01: 0.23889; 02: 0.23117; 03: 0.24917; 04: 0.24133; 05: 0.255; 06: 0.25215; 07: 0.23833; 08: 0.23833; 09: 0.24433; 10: 0.262; 11: 0.24333; 12: 0.20667
| Median |  | 01: 0.23922; 02: 0.2275; 03: 0.2495; 04: 0.242; 05: 0.2555; 06: 0.2527; 07: 0.24; 08: 0.24; 09: 0.241; 10: 0.262; 11: 0.24; 12: 0.205
| Std Dev. |  | 01: 0.007254; 02: 0.012545; 03: 0.003656; 04: 0.0048854; 05: 0.0052536; 06: 0.0072539; 07: 0.0040825; 08: 0.0040825; 09: 0.0082381; 10: 0.0062929; 11: 0.010328; 12: 0.008165
| Rel.Std.Dev. |  | 01: 3.037%; 02: 5.427%; 03: 1.467%; 04: 2.024%; 05: 2.06%; 06: 2.877%; 07: 1.713%; 08: 1.713%; 09: 3.372%; 10: 2.402%; 11: 4.244%; 12: 3.951%
| PDM3 |  | 01: -1.64%; 02: -4.819%; 03: 2.592%; 04: -0.6329%; 05: 4.994%; 06: 3.821%; 07: -1.868%; 08: -1.868%; 09: 0.6023%; 10: 7.876%; 11: 0.1906%; 12: -14.91%
| Z-Score (Absolute) |  | 01: 0.43; 02: 1.39; 03: 0.85; 04: 0.12; 05: 1.58; 06: 1.22; 07: 0.5; 08: 0.5; 09: 0.25; 10: 2.45; 11: 0.12; 12: 4.44
| Analytical results for B in OREAS 501b (Indicative Value 33.8 ppm) | [24 empty] | Indicative
B | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 02 | 04 | 05 | 06 | 07 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 25 | 27 | [8 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES | [8 empty] | 1
1
| 1 | 1 | < 2 | < 10 | < 10 | 4 | 50 | < 20 | < 10 | < 10 | 1.7351 | < 10 | < 10 | 76 | < 10 | 18 | 40 | [8 empty] | 1
| 1 | 2 | < 2 | < 10 | < 10 | 4 | 50 | < 20 | < 10 | < 10 | 1.7224 | < 10 | < 10 | 82 | < 10 | 18 | 39 | [8 empty] | 3
| 1 | 3 | < 2 | < 10 | < 10 | 4 | 60 | < 20 | < 10 | < 10 | 1.8224 | < 10 | < 10 | 93 | < 10 | 18 | 40 | [8 empty] | 16
| 1 | 4 | < 2 | < 10 | < 10 | 4 | 60 | < 20 | < 10 | < 10 | 1.955 | < 10 | < 10 | 111 | < 10 | 18 | 42 | [8 empty] | 33.8077
| 1 | 5 | < 2 | < 10 | < 10 | 4 | 60 | < 20 | < 10 | < 10 | 1.996 | < 10 | < 10 | 67 | < 10 | 18 | 42 | [8 empty] | 13
| 1 | 6 | < 2 | < 10 | < 10 | 3 | 60 | < 20 | < 10 | < 10 | 3.0194 | < 10 | < 10 | 62 | < 10 | 17 | 41
| Mean |  | 04: 3.8333; 05: 56.6667; 09: 2.0417; 12: 81.8333; 16: 17.8333; 17: 40.6667
| Median |  | 04: 4; 05: 60; 09: 1.8887; 12: 79; 16: 18; 17: 40.5
| Std Dev. |  | 04: 0.40825; 05: 5.164; 09: 0.49186; 12: 18.0157; 16: 0.40825; 17: 1.2111
| Rel.Std.Dev. |  | 04: 10.65%; 05: 9.113%; 09: 24.09%; 12: 22.02%; 16: 2.289%; 17: 2.978%
| PDM3 |  | 04: -88.66%; 05: 67.61%; 09: -93.96%; 12: 142.1%; 16: -47.25%; 17: 20.29%
| Z-Score (Absolute) |  | 0.87 | 0 | 0 | 0.67 | 5.1 | 1.09 | 0 | 0 | 0.87 | 0 | 0 | 7.85 | 0 | 0.86 | 3.35
| Analytical results for Ba in OREAS 501b (Certified Value 369 ppm) | [24 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  | 0
0
| 1 | 1 | 397.11 | 368 | 360 | 370 | 355 | 369 | 384 | 41 | 319 | 365 | 380 | 340 | 361.8571 | 370 | 370 | 312 | 390 | 384 | 313 | 370 | 358 | 380 |  | 1
| 1 | 2 | 399.78 | 352 | 360 | 360 | 351 | 382 | 394 | 39 | 321 | 372 | 380 | 350 | 363.6411 | 370 | 360 | 314 | 420 | 396 | 309 | 370 | 354 | 380 |  | #N/A
| 1 | 3 | 390.6 | 354 | 360 | 370 | 347 | 374 | 382 | 39 | 323 | 376 | 370 | 340 | 352.0344 | 380 | 370 | 349 | 400 | 391 | 314 | 362 | 382 | 380 |  | 16
| 1 | 4 | 398.64 | 372 | 360 | 370 | 357 | 383 | 382 | 39 | 320 | 380 | 380 | 330 | 366.5386 | 380 | 380 | 414 | 390 | 381 | 325 | 351 | 437 | 380 |  | 368.9409
| 1 | 5 | 391.8 | 360 | 360 | 360 | 357 | 384 | 385 | 41 | 321 | 372 | 380 | 330 | 361.2956 | 380 | 360 | 352 | 400 | 376 | 316 | 358 | 403 | 380 |  | 69
| 1 | 6 | 389.87 | 358 | 370 | 370 | 332 | 375 | 378 | 39 | 325 | 380 | 390 | 330 | 357.3461 | 390 | 360 | 330 | 400 | 393 | 321 | 348 | 407 | 400
| Mean |  | 394.6333 | 360.6667 | 361.6667 | 366.6667 | 349.8333 | 377.8333 | 384.1667 | 39.6667 | 321.5 | 374.1667 | 380 | 336.6667 | 360.4521 | 378.3333 | 366.6667 | 345.1667 | 400 | 386.8333 | 316.3333 | 359.8333 | 390.1667 | 383.3333
| Median |  | 394.455 | 359 | 360 | 370 | 353 | 378.5 | 383 | 39 | 321 | 374 | 380 | 335 | 361.5763 | 380 | 365 | 339.5 | 400 | 387.5 | 315 | 360 | 392.5 | 380
| Std Dev. |  | 4.374 | 7.8655 | 4.0825 | 5.164 | 9.5586 | 6.047 | 5.3821 | 1.0328 | 2.1679 | 5.7417 | 6.3246 | 8.165 | 5.107 | 7.5277 | 8.165 | 37.6851 | 10.9545 | 7.7309 | 5.785 | 9.3041 | 31.7831 | 8.165
| Rel.Std.Dev. |  | 1.108% | 2.181% | 1.129% | 1.408% | 2.732% | 1.6% | 1.401% | 2.604% | 0.6743% | 1.535% | 1.664% | 2.425% | 1.417% | 1.99% | 2.227% | 10.92% | 2.739% | 1.999% | 1.829% | 2.586% | 8.146% | 2.13%
| PDM3 |  | 6.964% | -2.243% | -1.972% | -0.6164% | -5.179% | 2.41% | 4.127% | -89.25% | -12.86% | 1.416% | 2.998% | -8.748% | -2.301% | 2.546% | -0.6164% | -6.444% | 8.418% | 4.85% | -14.26% | -2.469% | 5.753% | 3.901%
| Z-Score (Absolute) |  | 1.13 | 0.24 | 0.2 | 0 | 0.68 | 0.45 | 0.7 | 13.16 | 1.82 | 0.3 | 0.54 | 1.21 | 0.25 | 0.47 | 0 | 0.87 | 1.34 | 0.81 | 2.03 | 0.28 | 0.95 | 0.67
| Analytical results for Be in OREAS 501b (Certified Value 0.49 ppm) | [24 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 04 | 05 | 07 | 08 | 09 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [5 empty] | 2
2
| 1 | 1 | < 1 | 0.59 | 0.41 | 0.57 | 0.4 | 0.48 | < 0.5 | 0.5 | < 0.5 | 0.48 | 0.37 | 0.4 | < 2 | 0.48 | < 0.5 | 0.8 | 0.8 | < 5 | [5 empty] | 1
| 1 | 2 | < 1 | 0.54 | 0.43 | 0.61 | 0.4 | 0.48 | 0.5 | 0.5 | < 0.5 | 0.46 | 0.37 | 0.42 | < 2 | 0.52 | < 0.5 | 0.7 | 0.9 | < 5 | [5 empty] | #N/A
| 1 | 3 | < 1 | 0.58 | 0.38 | 0.59 | 0.5 | 0.48 | 0.5 | 0.5 | 0.5 | 0.52 | 0.39 | 0.38 | < 2 | 0.48 | < 0.5 | 0.8 | 0.8 | < 5 | [5 empty] | 16
| 1 | 4 | < 1 | 0.57 | 0.44 | 0.55 | 0.5 | 0.5 | < 0.5 | 0.5 | < 0.5 | 0.51 | 0.4 | 0.43 | < 2 | 0.51 | < 0.5 | 0.8 | 0.9 | < 5 | [5 empty] | 0.48939
| 1 | 5 | < 1 | 0.59 | 0.41 | 0.47 | 0.5 | 0.5 | < 0.5 | 0.5 | 0.5 | 0.51 | 0.47 | 0.43 | < 2 | 0.47 | < 0.5 | 0.7 | 0.8 | < 5 | [5 empty] | 70
| 1 | 6 | < 1 | 0.54 | 0.4 | 0.47 | 0.5 | 0.5 | < 0.5 | 0.5 | < 0.5 | 0.52 | 0.4 | 0.39 | < 2 | 0.51 | < 0.5 | 0.8 | 0.9 | < 5
| Mean |  |  | 0.56833 | 0.41167 | 0.54333 | 0.46667 | 0.49 | 0.5 | 0.5 | 0.5 | 0.5 | 0.4 | 0.40833 |  | 0.495 |  | 0.76667 | 0.85
| Median |  |  | 0.575 | 0.41 | 0.56 | 0.5 | 0.49 | 0.5 | 0.5 | 0.5 | 0.51 | 0.395 | 0.41 |  | 0.495 |  | 0.8 | 0.85
| Std Dev. |  |  | 0.023166 | 0.02137 | 0.060222 | 0.05164 | 0.010954 | 0 | 0 | 0 | 0.024495 | 0.036878 | 0.02137 |  | 0.020736 |  | 0.05164 | 0.054772
| Rel.Std.Dev. |  |  | 4.076% | 5.191% | 11.08% | 11.07% | 2.236% | 0% | 0% | 0% | 4.899% | 9.22% | 5.233% |  | 4.189% |  | 6.736% | 6.444%
| PDM3 |  |  | 16.13% | -15.88% | 11.02% | -4.644% | 0.1238% | 2.167% | 2.167% | 2.167% | 2.167% | -18.27% | -16.56% |  | 1.146% |  | 56.66% | 73.68%
| Z-Score (Absolute) |  | 6.04 | 0.83 | 1.07 | 0.52 | 0.4 | 0.12 | 0 | 0 | 0 | 0 | 1.21 | 1.11 | 18.12 | 0.06 | 0 | 3.22 | 4.23 | 54.35
| Analytical results for Bi in OREAS 501b (Certified Value 1.6 ppm) | [24 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  | 2
3
| 1 | 1 | 1.443 | 1.56 | 1.5 | 1.26 | 1.48 | 1.7 | 1.61 | 1.56 | < 10 | < 2 | < 3 | 1.79 | 2 | 1.618 | 1.46 | 1.63 | 9 | 1.47 | 1.6 | 1.77 | < 1 | < 2 |  | 1
| 1 | 2 | 1.4942 | 1.58 | 1.5 | 1.92 | 1.71 | 1.9 | 1.59 | 1.58 | < 10 | < 2 | < 3 | 2.06 | 5 | 1.642 | 1.37 | 1.67 | 1 | 1.55 | 1.7 | 1.7 | < 1 | < 2 |  | #N/A
| 1 | 3 | 1.4576 | 1.56 | 1.56 | 1.61 | 1.5 | 1.6 | 1.63 | 1.56 | < 10 | < 2 | < 3 | 1.76 | 3 | 1.704 | 1.46 | 1.65 | 2 | 1.59 | 1.7 | 1.69 | < 1 | < 2 |  | 16
| 1 | 4 | 1.4269 | 1.6 | 1.58 | 1.62 | 1.52 | 1.6 | 1.8 | 1.61 | < 10 | < 2 | < 3 | 1.83 | 6 | 1.797 | 1.37 | 1.66 | < 1 | 1.59 | 1.7 | 1.68 | < 1 | < 2 |  | 1.603
| 1 | 5 | 1.4279 | 1.54 | 1.5 | 1.63 | 1.86 | 1.8 | 1.64 | 1.59 | < 10 | < 2 | < 3 | 1.73 | 3 | 1.679 | 1.46 | 1.59 | 18 | 1.56 | 1.7 | 1.8 | < 1 | < 2 |  | 71
| 1 | 6 | 1.4716 | 1.55 | 1.48 | 1.56 | 1.51 | 1.6 | 1.62 | 1.49 | < 10 | < 2 | < 3 | 1.84 | 5 | 1.671 | 1.48 | 1.5 | 1 | 1.53 | 1.7 | 1.7 | < 1 | < 2
| Mean |  | 1.4535 | 1.565 | 1.52 | 1.6 | 1.5967 | 1.7 | 1.6483 | 1.565 |  |  |  | 1.835 | 4 | 1.6852 | 1.4333 | 1.6167 | 6.2 | 1.5483 | 1.6833 | 1.7233
| Median |  | 1.4503 | 1.56 | 1.5 | 1.615 | 1.515 | 1.65 | 1.625 | 1.57 |  |  |  | 1.81 | 4 | 1.675 | 1.46 | 1.64 | 2 | 1.555 | 1.7 | 1.7
| Std Dev. |  | 0.026359 | 0.021679 | 0.04 | 0.21024 | 0.15397 | 0.12649 | 0.076267 | 0.041352 |  |  |  | 0.11777 | 1.5492 | 0.06241 | 0.049666 | 0.06377 | 7.3959 | 0.044907 | 0.040825 | 0.049261
| Rel.Std.Dev. |  | 1.813% | 1.385% | 2.632% | 13.14% | 9.643% | 7.441% | 4.627% | 2.642% |  |  |  | 6.418% | 38.73% | 3.703% | 3.465% | 3.945% | 119.3% | 2.9% | 2.425% | 2.858%
| PDM3 |  | -9.327% | -2.373% | -5.181% | -0.1901% | -0.398% | 6.048% | 2.825% | -2.373% |  |  |  | 14.47% | 149.5% | 5.123% | -10.59% | 0.8496% | 286.8% | -3.413% | 5.008% | 7.504%
| Z-Score (Absolute) |  | 1.09 | 0.52 | 0.75 | 0.34 | 0.35 | 0.17 | 0.09 | 0.52 | 42.68 | 1.71 | 6.83 | 0.87 | 11.95 | 0.1 | 1.19 | 0.25 | 18.78 | 0.6 | 0.09 | 0.29 | 3.41 | 1.71
| Analytical results for Ca in OREAS 501b (Certified Value 0.966 wt.%) | [24 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 3
3
| 1 | 1 | 1.1568 | 1.23 | 0.94 | 0.96 | 0.97 | 0.87 | 0.91 | 1.01 | 3.59 | 0.82 | 0.91 | 1.01 | 0.87 | 1.0076 | 0.93 | 0.99 | 0.7 | 0.99 | 1.03 | 0.95 | 1.01 | 1.34 | 1.04 | 1
| 1 | 2 | 1.1565 | 1.23 | 0.93 | 0.97 | 0.95 | 0.89 | 0.91 | 1.07 | 3.59 | 0.84 | 0.92 | 1 | 0.93 | 1.0477 | 0.95 | 0.96 | 0.71 | 0.99 | 1.07 | 0.96 | 0.97 | 1.27 | 1.03 | #N/A
| 1 | 3 | 1.1227 | 1.23 | 0.96 | 0.93 | 0.96 | 0.88 | 0.89 | 1.04 | 3.58 | 0.86 | 0.91 | 1.01 | 0.92 | 1.0008 | 0.98 | 0.98 | 0.78 | 1 | 1.04 | 1 | 0.98 | 1.31 | 1.04 | 16
| 1 | 4 | 1.1209 | 1.27 | 0.97 | 0.95 | 0.95 | 0.86 | 0.91 | 1.05 | 3.65 | 0.85 | 0.92 | 1.03 | 0.9 | 1.021 | 0.96 | 0.99 | 0.9 | 1.01 | 1.03 | 1.01 | 0.97 | 1.4 | 1.04 | 0.96588
| 1 | 5 | 1.0975 | 1.24 | 0.95 | 0.96 | 0.95 | 0.87 | 0.9 | 1.07 | 3.59 | 0.87 | 0.89 | 1.03 | 0.91 | 1.0222 | 0.98 | 0.97 | 0.8 | 0.99 | 1.01 | 0.98 | 0.98 | 1.23 | 1.04 | 72
| 1 | 6 | 1.1491 | 1.22 | 0.96 | 0.96 | 0.95 | 0.88 | 0.91 | 1.04 | 3.51 | 0.85 | 0.92 | 1.02 | 0.9 | 1.0135 | 0.99 | 0.95 | 0.75 | 1 | 1.06 | 1.04 | 0.97 | 1.42 | 1.07
| Mean |  | 1.1339 | 1.2367 | 0.95167 | 0.955 | 0.955 | 0.875 | 0.905 | 1.0467 | 3.585 | 0.84833 | 0.91167 | 1.0167 | 0.905 | 1.0188 | 0.965 | 0.97333 | 0.77333 | 0.99667 | 1.04 | 0.99 | 0.98 | 1.3283 | 1.0433
| Median |  | 1.1359 | 1.23 | 0.955 | 0.96 | 0.95 | 0.875 | 0.91 | 1.045 | 3.59 | 0.85 | 0.915 | 1.015 | 0.905 | 1.0173 | 0.97 | 0.975 | 0.765 | 0.995 | 1.035 | 0.99 | 0.975 | 1.325 | 1.04
| Std Dev. |  | 0.024023 | 0.017512 | 0.01472 | 0.013784 | 0.0083666 | 0.010488 | 0.0083666 | 0.022509 | 0.044609 | 0.017224 | 0.01169 | 0.012111 | 0.020736 | 0.016308 | 0.022583 | 0.01633 | 0.073121 | 0.008165 | 0.021909 | 0.033466 | 0.015492 | 0.073598 | 0.013663
| Rel.Std.Dev. |  | 2.119% | 1.416% | 1.547% | 1.443% | 0.8761% | 1.199% | 0.9245% | 2.151% | 1.244% | 2.03% | 1.282% | 1.191% | 2.291% | 1.601% | 2.34% | 1.678% | 9.455% | 0.8192% | 2.107% | 3.38% | 1.581% | 5.541% | 1.31%
| PDM3 |  | 17.4% | 28.04% | -1.471% | -1.126% | -1.126% | -9.409% | -6.303% | 8.364% | 271.2% | -12.17% | -5.613% | 5.258% | -6.303% | 5.479% | -0.09084% | 0.7719% | -19.93% | 3.188% | 7.674% | 2.497% | 1.462% | 37.53% | 8.019%
| Z-Score (Absolute) |  | 1.64 | 2.73 | 0.3 | 0.27 | 0.27 | 1.12 | 0.8 | 0.71 | 27.74 | 1.4 | 0.73 | 0.39 | 0.8 | 0.41 | 0.16 | 0.07 | 2.2 | 0.18 | 0.64 | 0.11 | 0 | 3.71 | 0.67
| Analytical results for Cd in OREAS 501b (Indicative Value 0.1 ppm) | [24 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  | 2
2
| 1 | 1 | 0.2085 | 0.08 | 0.05 | 0.09 | 0.08 | 0.2 | 0.13 | 0.08 | < 1 | < 0.2 | < 0.5 | 0.12 | < 0.5 | 0.086957 | 0.14 | 0.09 | < 4 | 0.12 | < 0.2 | 0.03 | 4 | 0.59 |  | 1
| 1 | 2 | 0.2032 | 0.09 | < 0.05 | 0.09 | 0.1 | 0.2 | 0.13 | 0.09 | < 1 | < 0.2 | < 0.5 | 0.1 | < 0.5 | 0.10373 | 0.13 | 0.09 | < 4 | 0.1 | < 0.2 | 0.06 | 3 | 0.59 |  | 4
| 1 | 3 | 0.21 | 0.09 | < 0.05 | 0.08 | 0.12 | 0.1 | 0.15 | 0.1 | < 1 | < 0.2 | < 0.5 | 0.1 | < 0.5 | 0.1 | 0.13 | 0.09 | < 4 | 0.08 | < 0.2 | 0.06 | 4 | 0.57 |  | 16
| 1 | 4 | 0.1948 | 0.09 | < 0.05 | 0.08 | 0.09 | 0.1 | 0.13 | 0.12 | < 1 | < 0.2 | < 0.5 | 0.09 | < 0.5 | 0.1084 | 0.13 | 0.07 | < 4 | 0.1 | < 0.2 | 0.04 | 5 | 0.52 |  | 0.1011
| 1 | 5 | 0.2033 | 0.1 | < 0.05 | 0.08 | 0.15 | 0.1 | 0.12 | 0.12 | < 1 | < 0.2 | < 0.5 | 0.11 | < 0.5 | 0.085062 | 0.13 | 0.07 | < 4 | 0.07 | < 0.2 | 0.06 | 4 | < 0.5 |  | 14
| 1 | 6 | 0.2114 | 0.06 | < 0.05 | 0.08 | 0.11 | 0.2 | 0.12 | 0.09 | < 1 | < 0.2 | < 0.5 | 0.1 | < 0.5 | 0.09188 | 0.14 | 0.06 | < 4 | 0.06 | < 0.2 | 0.05 | 5 | 0.57
| Mean |  | 0.2052 | 0.085 | 0.05 | 0.083333 | 0.10833 | 0.15 | 0.13 | 0.1 |  |  |  | 0.10333 |  | 0.096005 | 0.13333 | 0.078333 |  | 0.088333 |  | 0.05 | 4.1667 | 0.568
| Median |  | 0.2059 | 0.09 | 0.05 | 0.08 | 0.105 | 0.15 | 0.13 | 0.095 |  |  |  | 0.1 |  | 0.09594 | 0.13 | 0.08 |  | 0.09 |  | 0.055 | 4 | 0.57
| Std Dev. |  | 0.0061342 | 0.013784 |  | 0.005164 | 0.024833 | 0.054772 | 0.010954 | 0.016733 |  |  |  | 0.010328 |  | 0.009465 | 0.005164 | 0.013292 |  | 0.022286 |  | 0.012649 | 0.75277 | 0.028636
| Rel.Std.Dev. |  | 2.989% | 16.22% |  | 6.197% | 22.92% | 36.51% | 8.427% | 16.73% |  |  |  | 9.995% |  | 9.859% | 3.873% | 16.97% |  | 25.23% |  | 25.3% | 18.07% | 5.041%
| PDM3 |  | 103% | -15.93% | -50.54% | -17.57% | 7.154% | 48.37% | 28.58% | -1.089% |  |  |  | 2.208% |  | -5.04% | 31.88% | -22.52% |  | -12.63% |  | -50.54% | 4021% | 461.8%
| Z-Score (Absolute) |  | 0.82 | 0.52 | 0.91 | 0.54 | 0.26 | 0.2 | 0.02 | 0.35 | 9.63 | 0.76 | 4.08 | 0.31 | 4.08 | 0.4 | 0.02 | 0.59 | 42.88 | 0.48 | 0.76 | 0.91 | 44.73 | 4.71
| Analytical results for Ce in OREAS 501b (Certified Value 61 ppm) | [24 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 0
1
| 1 | 1 | 61.347 | 63.1 | 59.6 | 59.7 | 59.2 | 59.4 | 56.58 | 60.6 | 53.041 | 66.7 | 64.6 | 36 | 52.7 | 60.2 | 62.57 | [8 empty] | 1
| 1 | 2 | 63.369 | 62.3 | 57.8 | 58.6 | 63.2 | 57.2 | 61.82 | 59.5 | 52.103 | 68.6 | 64 | 34 | 61.5 | 61 | 50.23 | [8 empty] | #N/A
| 1 | 3 | 61.858 | 64.28 | 60.1 | 56.2 | 64.4 | 55.5 | 58.98 | 60.2 | 54.079 | 71.2 | 62.5 | 38 | 60.4 | 63 | 51.97 | [8 empty] | 16
| 1 | 4 | 60.177 | 60.38 | 60.7 | 58.4 | 64.7 | 56.1 | 60.13 | 60.8 | 57.083 | 70 | 65.7 | 50 | 60.6 | 64.9 | 48.62 | [8 empty] | 60.5096
| 1 | 5 | 60.476 | 62.03 | 58.9 | 55.7 | 62.3 | 55.4 | 61.67 | 60.6 | 53.016 | 61.5 | 64.7 | 35 | 59.5 | 64.7 | 62.73 | [8 empty] | 73
| 1 | 6 | 63.426 | 61.69 | 57.3 | 58.9 | 60.8 | 57.5 | 58.11 | 60.5 | 52.13 | 68.9 | 59.6 | 35 | 64.2 | 60.3 | 57.75
| Mean |  | 61.7755 | 62.2967 | 59.0667 | 57.9167 | 62.4333 | 56.85 | 59.5483 | 60.3667 | 53.5753 | 67.8167 | 63.5167 | 38 | 59.8167 | 62.35 | 55.645
| Median |  | 61.6025 | 62.165 | 59.25 | 58.5 | 62.75 | 56.65 | 59.555 | 60.55 | 53.0285 | 68.75 | 64.3 | 35.5 | 60.5 | 62 | 54.86
| Std Dev. |  | 1.3927 | 1.3185 | 1.3246 | 1.5943 | 2.1323 | 1.5189 | 2.0593 | 0.46762 | 1.8661 | 3.4394 | 2.1885 | 6.0332 | 3.8405 | 2.1492 | 6.2412
| Rel.Std.Dev. |  | 2.254% | 2.116% | 2.243% | 2.753% | 3.415% | 2.672% | 3.458% | 0.7746% | 3.483% | 5.072% | 3.446% | 15.88% | 6.421% | 3.447% | 11.22%
| PDM3 |  | 2.092% | 2.953% | -2.385% | -4.285% | 3.179% | -6.048% | -1.589% | -0.2363% | -11.46% | 12.08% | 4.97% | -37.2% | -1.145% | 3.041% | -8.039%
| Z-Score (Absolute) |  | 0.52 | 0.66 | 0.2 | 0.51 | 0.7 | 0.79 | 0.07 | 0.15 | 1.66 | 2.13 | 0.98 | 5.81 | 0 | 0.67 | 1.11
| Analytical results for Co in OREAS 501b (Certified Value 14.9 ppm) | [24 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  | 1
2
| 1 | 1 | 14.914 | 15.1 | 15 | 14.7 | 15.3 | 14.2 | 15.1 | 15.4 | 10 | 15 | 13 | 14.4 | 14 | 15.188 | 12.6 | 15.4 | 15 | 14.2 | 14.8 | 16.6 | 16 | 23 |  | 1
| 1 | 2 | 15.576 | 15.2 | 17 | 15.1 | 15.7 | 14.4 | 15.5 | 14.9 | 10 | 15 | 13 | 13.8 | 13 | 14.9222 | 12.9 | 16.3 | 15 | 14.8 | 15.1 | 16.8 | 14 | 21 |  | #N/A
| 1 | 3 | 15.821 | 14.9 | 16 | 14.3 | 16.2 | 13.6 | 15.6 | 14.4 | 9 | 15 | 12 | 14.3 | 15 | 15.1719 | 13 | 15.3 | 17 | 14.6 | 15.2 | 16 | 18 | 20 |  | 16
| 1 | 4 | 15.647 | 14.8 | 17 | 14.9 | 15.2 | 14.5 | 15.7 | 14.6 | 10 | 15 | 13 | 13.7 | 13 | 15.5576 | 13.1 | 16.1 | 19 | 14.8 | 15.2 | 16.1 | 19 | 23 |  | 14.8524
| 1 | 5 | 15.844 | 15.2 | 16 | 14.9 | 15.5 | 14.6 | 15.4 | 14.9 | 10 | 14 | 13 | 13.7 | 13 | 15.7822 | 14 | 15.4 | 16 | 14.6 | 15 | 16 | 18 | 23 |  | 74
| 1 | 6 | 15.776 | 14.8 | 18 | 15.3 | 15.3 | 14.4 | 15 | 14.7 | 9 | 14 | 13 | 14.1 | 13 | 14.9893 | 13.1 | 14.7 | 15 | 15.1 | 15 | 16.3 | 19 | 22
| Mean |  | 15.5963 | 15 | 16.5 | 14.8667 | 15.5333 | 14.2833 | 15.3833 | 14.8167 | 9.6667 | 14.6667 | 12.8333 | 14 | 13.5 | 15.2685 | 13.1167 | 15.5333 | 16.1667 | 14.6833 | 15.05 | 16.3 | 17.3333 | 22
| Median |  | 15.7115 | 15 | 16.5 | 14.9 | 15.4 | 14.4 | 15.45 | 14.8 | 10 | 15 | 13 | 13.95 | 13 | 15.18 | 13.05 | 15.4 | 15.5 | 14.7 | 15.05 | 16.2 | 18 | 22.5
| Std Dev. |  | 0.35004 | 0.18974 | 1.0488 | 0.34448 | 0.37238 | 0.36009 | 0.27869 | 0.34303 | 0.5164 | 0.5164 | 0.40825 | 0.30984 | 0.83666 | 0.335 | 0.47081 | 0.58195 | 1.6021 | 0.29944 | 0.15166 | 0.33466 | 1.9664 | 1.2649
| Rel.Std.Dev. |  | 2.244% | 1.265% | 6.356% | 2.317% | 2.397% | 2.521% | 1.812% | 2.315% | 5.342% | 3.521% | 3.181% | 2.213% | 6.197% | 2.194% | 3.589% | 3.746% | 9.91% | 2.039% | 1.008% | 2.053% | 11.34% | 5.75%
| PDM3 |  | 5.009% | 0.9941% | 11.09% | 0.09639% | 4.585% | -3.831% | 3.575% | -0.2403% | -34.91% | -1.25% | -13.59% | -5.739% | -9.105% | 2.802% | -11.69% | 4.585% | 8.849% | -1.138% | 1.331% | 9.747% | 16.7% | 48.12%
| Z-Score (Absolute) |  | 0.59 | 0.03 | 1.52 | 0.16 | 0.52 | 0.76 | 0.37 | 0.21 | 5.5 | 0.37 | 2.25 | 1.05 | 1.57 | 0.25 | 1.96 | 0.52 | 1.17 | 0.35 | 0.03 | 1.31 | 2.37 | 7.16
| Analytical results for Cr in OREAS 501b (Certified Value 85 ppm) | [24 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 0
1
| 1 | 1 | 95.511 | 91 | 85 | 85 | 81 | 87 | 84 | 87 | 21 | 83 | 84 | 85 | 83 | 84.3902 | 82 | 86 | 73 | 86 | 89 | 88 | 76 | 67 | 129 | 1
| 1 | 2 | 94.989 | 91 | 85 | 85 | 87 | 87 | 87 | 91 | 21 | 83 | 83 | 87 | 97 | 82.3154 | 82 | 83 | 75 | 88 | 84 | 86 | 79 | 65 | 125 | #N/A
| 1 | 3 | 92.49 | 93 | 85 | 79 | 91 | 87 | 87 | 89 | 21 | 84 | 85 | 84 | 83 | 84.5896 | 82 | 86 | 83 | 86 | 87 | 88 | 77 | 75 | 126 | 16
| 1 | 4 | 93.704 | 93 | 85 | 82 | 82 | 86 | 86 | 88 | 21 | 96 | 85 | 86 | 84 | 86.1924 | 81 | 84 | 98 | 86 | 88 | 90 | 79 | 85 | 132 | 85.2696
| 1 | 5 | 90.9 | 93 | 85 | 83 | 80 | 88 | 85 | 91 | 21 | 84 | 84 | 86 | 79 | 87.8776 | 83 | 82 | 86 | 85 | 87 | 85 | 75 | 77 | 128 | 75
| 1 | 6 | 93.206 | 91 | 80 | 84 | 81 | 88 | 85 | 88 | 21 | 82 | 87 | 88 | 82 | 82.5865 | 89 | 82 | 77 | 88 | 92 | 89 | 78 | 80 | 131
| Mean |  | 93.4667 | 92 | 84.1667 | 83 | 83.6667 | 87.1667 | 85.6667 | 89 | 21 | 85.3333 | 84.6667 | 86 | 84.6667 | 84.6586 | 83.1667 | 83.8333 | 82 | 86.5 | 87.8333 | 87.6667 | 77.3333 | 74.8333 | 128.5
| Median |  | 93.455 | 92 | 85 | 83.5 | 81.5 | 87 | 85.5 | 88.5 | 21 | 83.5 | 84.5 | 86 | 83 | 84.4899 | 82 | 83.5 | 80 | 86 | 87.5 | 88 | 77.5 | 76 | 128.5
| Std Dev. |  | 1.683 | 1.0954 | 2.0412 | 2.2804 | 4.3665 | 0.75277 | 1.2111 | 1.6733 | 0 | 5.2789 | 1.3663 | 1.4142 | 6.2823 | 2.1249 | 2.9269 | 1.8348 | 9.252 | 1.2247 | 2.6394 | 1.8619 | 1.633 | 7.6529 | 2.7386
| Rel.Std.Dev. |  | 1.801% | 1.191% | 2.425% | 2.747% | 5.219% | 0.8636% | 1.414% | 1.88% | 0% | 6.186% | 1.614% | 1.644% | 7.42% | 2.51% | 3.519% | 2.189% | 11.28% | 1.416% | 3.005% | 2.124% | 2.112% | 10.23% | 2.131%
| PDM3 |  | 9.613% | 7.893% | -1.293% | -2.662% | -1.88% | 2.225% | 0.4657% | 4.375% | -75.37% | 0.07475% | -0.7071% | 0.8566% | -0.7071% | -0.7165% | -2.466% | -1.684% | -3.834% | 1.443% | 3.007% | 2.811% | -9.307% | -12.24% | 50.7%
| Z-Score (Absolute) |  | 3.24 | 2.7 | 0.18 | 0.61 | 0.37 | 0.92 | 0.37 | 1.59 | 23.42 | 0.25 | 0 | 0.49 | 0 | 0 | 0.55 | 0.31 | 0.98 | 0.67 | 1.16 | 1.1 | 2.7 | 3.62 | 16.12
| Analytical results for Cs in OREAS 501b (Certified Value 9.93 ppm) | [24 empty] | Certified
Cs | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 12 | 16 | 17 | 18 | 20 | 23 | 25 | [9 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
3
| 1 | 1 | 9.9622 | 9.46 | 9.86 | 9.51 | 9.66 | 9.21 | 9.7 | 9.75 | 10.4978 | 11.05 | 10.55 | 9.51 | 10.3 | 8.54 | [9 empty] | 1
| 1 | 2 | 10.339 | 9.22 | 9.58 | 9.65 | 10.5 | 9.56 | 9.96 | 9.52 | 10.611 | 11.35 | 10.65 | 9.47 | 10.5 | 8.26 | [9 empty] | #N/A
| 1 | 3 | 10.226 | 9.29 | 9.92 | 9.55 | 10.4 | 9.59 | 9.91 | 9.53 | 10.7646 | 11.3 | 10.3 | 9.54 | 10.5 | 8.32 | [9 empty] | 16
| 1 | 4 | 9.883 | 9.68 | 10.2 | 9.46 | 9.98 | 9.72 | 9.88 | 9.57 | 10.7061 | 11.15 | 10.4 | 9.8 | 10.7 | 8.27 | [9 empty] | 9.9256
| 1 | 5 | 10.038 | 9.31 | 9.56 | 9.56 | 10.25 | 9.62 | 9.8 | 9.53 | 10.7158 | 10.5 | 10.35 | 9.62 | 10.5 | 8.41 | [9 empty] | 76
| 1 | 6 | 10.344 | 9.24 | 9.68 | 9.78 | 10 | 9.71 | 9.71 | 9.7 | 10.4797 | 11.55 | 9.81 | 10.05 | 10.2 | 8.27
| Mean |  | 10.132 | 9.3667 | 9.8 | 9.585 | 10.1317 | 9.5683 | 9.8267 | 9.6 | 10.6292 | 11.15 | 10.3433 | 9.665 | 10.45 | 8.345
| Median |  | 10.132 | 9.3 | 9.77 | 9.555 | 10.125 | 9.605 | 9.84 | 9.55 | 10.6585 | 11.225 | 10.375 | 9.58 | 10.5 | 8.295
| Std Dev. |  | 0.19814 | 0.17523 | 0.24429 | 0.11432 | 0.31128 | 0.18691 | 0.10764 | 0.099599 | 0.11972 | 0.36194 | 0.292 | 0.22188 | 0.17607 | 0.11077
| Rel.Std.Dev. |  | 1.956% | 1.871% | 2.493% | 1.193% | 3.072% | 1.953% | 1.095% | 1.037% | 1.126% | 3.246% | 2.823% | 2.296% | 1.685% | 1.327%
| PDM3 |  | 2.08% | -5.631% | -1.265% | -3.431% | 2.076% | -3.599% | -0.9965% | -3.28% | 7.089% | 12.34% | 4.209% | -2.625% | 5.284% | -15.92%
| Z-Score (Absolute) |  | 0.67 | 0.95 | 0.03 | 0.48 | 0.67 | 0.52 | 0.03 | 0.45 | 1.73 | 2.83 | 1.12 | 0.31 | 1.35 | 3.11
| Analytical results for Cu in OREAS 501b (Certified Value 0.258 wt.%) | [24 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 3
3
| 1 | 1 | 0.28838 | 0.26709 | 0.252 | 0.257 | 0.266 | 0.25033 | 0.247 | 0.25967 | 0.0062 | 0.243 | 0.2572 | 0.264 | 0.255 | 0.25612 | 0.249 | 0.272 | 0.225 | 0.258 | 0.2507 | 0.2483 | 0.29385 | 0.23 | 0.2582 | 1
| 1 | 2 | 0.29111 | 0.2682 | 0.253 | 0.263 | 0.261 | 0.25971 | 0.253 | 0.26437 | 0.0062 | 0.247 | 0.2637 | 0.261 | 0.262 | 0.26174 | 0.254 | 0.268 | 0.2272 | 0.262 | 0.253 | 0.2502 | 0.29392 | 0.2 | 0.2563 | #N/A
| 1 | 3 | 0.28051 | 0.26771 | 0.259 | 0.252 | 0.262 | 0.25278 | 0.247 | 0.25605 | 0.0062 | 0.249 | 0.2585 | 0.265 | 0.263 | 0.25742 | 0.258 | 0.269 | 0.2504 | 0.262 | 0.2544 | 0.2581 | 0.292 | 0.23 | 0.259 | 16
| 1 | 4 | 0.28187 | 0.25928 | 0.269 | 0.258 | 0.268 | 0.25276 | 0.253 | 0.25776 | 0.0062 | 0.243 | 0.2634 | 0.268 | 0.252 | 0.26689 | 0.254 | 0.274 | 0.2658 | 0.265 | 0.2551 | 0.2614 | 0.2911 | 0.24 | 0.2606 | 0.25812
| 1 | 5 | 0.27408 | 0.26148 | 0.251 | 0.259 | 0.263 | 0.25393 | 0.251 | 0.26035 | 0.0062 | 0.249 | 0.2556 | 0.269 | 0.251 | 0.26539 | 0.257 | 0.266 | 0.2501 | 0.26 | 0.2558 | 0.2542 | 0.29472 | 0.25 | 0.2633 | 77
| 1 | 6 | 0.28431 | 0.2653 | 0.26 | 0.26 | 0.261 | 0.25379 | 0.252 | 0.25109 | 0.0062 | 0.246 | 0.268 | 0.265 | 0.256 | 0.25493 | 0.262 | 0.27 | 0.2351 | 0.261 | 0.2504 | 0.2696 | 0.28813 | 0.25 | 0.2685
| Mean |  | 0.28338 | 0.26484 | 0.25733 | 0.25817 | 0.2635 | 0.25388 | 0.2505 | 0.25821 | 0.0062 | 0.24617 | 0.26107 | 0.26533 | 0.2565 | 0.26042 | 0.25567 | 0.26983 | 0.24227 | 0.26133 | 0.25323 | 0.25697 | 0.29229 | 0.23333 | 0.26098
| Median |  | 0.28309 | 0.2662 | 0.256 | 0.2585 | 0.2625 | 0.25328 | 0.2515 | 0.25871 | 0.0062 | 0.2465 | 0.26095 | 0.265 | 0.2555 | 0.25958 | 0.2555 | 0.2695 | 0.2426 | 0.2615 | 0.2537 | 0.25615 | 0.29292 | 0.235 | 0.2598
| Std Dev. |  | 0.0060436 | 0.0036607 | 0.0068313 | 0.003656 | 0.002881 | 0.0031326 | 0.0028107 | 0.004477 | 0 | 0.0027142 | 0.0047302 | 0.0028752 | 0.00501 | 0.0050223 | 0.0044121 | 0.0028577 | 0.015861 | 0.0023381 | 0.0022774 | 0.0078625 | 0.0024414 | 0.018619 | 0.0043733
| Rel.Std.Dev. |  | 2.133% | 1.382% | 2.655% | 1.416% | 1.093% | 1.234% | 1.122% | 1.734% | 0% | 1.103% | 1.812% | 1.084% | 1.953% | 1.929% | 1.726% | 1.059% | 6.547% | 0.8947% | 0.8993% | 3.06% | 0.8353% | 7.98% | 1.676%
| PDM3 |  | 9.785% | 2.605% | -0.305% | 0.01787% | 2.084% | -1.642% | -2.952% | 0.03659% | -97.6% | -4.631% | 1.141% | 2.794% | -0.6278% | 0.8891% | -0.9507% | 4.538% | -6.142% | 1.245% | -1.893% | -0.447% | 13.24% | -9.603% | 1.109%
| Z-Score (Absolute) |  | 3.45 | 0.91 | 0.11 | 0 | 0.73 | 0.59 | 1.05 | 0.01 | 34.44 | 1.64 | 0.4 | 0.98 | 0.23 | 0.31 | 0.34 | 1.59 | 2.17 | 0.43 | 0.67 | 0.16 | 4.66 | 3.39 | 0.38
| Analytical results for Dy in OREAS 501b (Indicative Value 3.22 ppm) | [24 empty] | Indicative
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS | 2
2
| 1 | 1 | 01: < 10; 02: 3.43; 03: 2.94; 04: 3.2; 05: 3.1 | 1
| 1 | 2 | 01: < 10; 02: 2.38; 03: 2.91; 04: 3.32; 05: 3.1 | 9
| 1 | 3 | 01: < 10; 02: 3.48; 03: 3.06; 04: 3.25; 05: 3.1 | 16
| 1 | 4 | 01: < 10; 02: 3.4; 03: 3.1; 04: 3.34; 05: 3.2 | 3.2221
| 1 | 5 | 01: < 10; 02: 3.61; 03: 2.92; 04: 3.4; 05: 3.2 | 15
| 1 | 6 | 01: < 10; 02: 3.58; 03: 2.89; 04: 3.2; 05: 3.1
| Mean |  | 02: 3.3133; 03: 2.97; 04: 3.285; 05: 3.1333
| Median |  | 02: 3.455; 03: 2.93; 04: 3.285; 05: 3.1
| Std Dev. |  | 02: 0.46457; 03: 0.087636; 04: 0.081425; 05: 0.05164
| Rel.Std.Dev. |  | 02: 14.02%; 03: 2.951%; 04: 2.479%; 05: 1.648%
| PDM3 |  | 02: 2.832%; 03: -7.824%; 04: 1.953%; 05: -2.754%
| Z-Score (Absolute) |  | 01: 29.85; 02: 0.13; 03: 1.4; 04: 0; 05: 0.67
| Analytical results for Er in OREAS 501b (Indicative Value 1.7 ppm) | [24 empty] | Indicative
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS | 2
2
| 1 | 1 | 01: < 10; 02: 1.6; 03: 1.56; 04: 1.75; 05: 1.7 | 1
| 1 | 2 | 01: < 10; 02: 1.13; 03: 1.63; 04: 1.83; 05: 1.7 | 10
| 1 | 3 | 01: < 10; 02: 1.69; 03: 1.66; 04: 1.78; 05: 1.7 | 16
| 1 | 4 | 01: < 10; 02: 1.65; 03: 1.68; 04: 1.76; 05: 1.7 | 1.6983
| 1 | 5 | 01: < 10; 02: 1.74; 03: 1.58; 04: 1.83; 05: 1.7 | 16
| 1 | 6 | 01: < 10; 02: 1.71; 03: 1.57; 04: 1.76; 05: 1.8
| Mean |  | 02: 1.5867; 03: 1.6133; 04: 1.785; 05: 1.7167
| Median |  | 02: 1.67; 03: 1.605; 04: 1.77; 05: 1.7
| Std Dev. |  | 02: 0.22897; 03: 0.050465; 04: 0.036194; 05: 0.040825
| Rel.Std.Dev. |  | 02: 14.43%; 03: 3.128%; 04: 2.028%; 05: 2.378%
| PDM3 |  | 02: -6.57%; 03: -5%; 04: 5.108%; 05: 1.084%
| Z-Score (Absolute) |  | 01: 54.05; 02: 0.85; 03: 0.67; 04: 0.45; 05: 0
| Analytical results for Eu in OREAS 501b (Indicative Value 0.38 ppm) | [24 empty] | Indicative
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS | 2
2
| 1 | 1 | 01: < 2; 02: 0.35; 03: 0.34; 04: 0.41; 05: 0.4 | 1
| 1 | 2 | 01: < 2; 02: 0.26; 03: 0.34; 04: 0.46; 05: 0.4 | 11
| 1 | 3 | 01: < 2; 02: 0.37; 03: 0.358; 04: 0.41; 05: 0.4 | 16
| 1 | 4 | 01: < 2; 02: 0.37; 03: 0.35; 04: 0.41; 05: 0.4 | 0.38242
| 1 | 5 | 01: < 2; 02: 0.37; 03: 0.332; 04: 0.42; 05: 0.4 | 17
| 1 | 6 | 01: < 2; 02: 0.36; 03: 0.334; 04: 0.43; 05: 0.4
| Mean |  | 02: 0.34667; 03: 0.34233; 04: 0.42333; 05: 0.4
| Median |  | 02: 0.365; 03: 0.34; 04: 0.415; 05: 0.4
| Std Dev. |  | 02: 0.043205; 03: 0.009913; 04: 0.019664; 05: 6.0809e-17
| Rel.Std.Dev. |  | 02: 12.46%; 03: 2.896%; 04: 4.645%; 05: 1.52e-14%
| PDM3 |  | 02: -9.348%; 03: -10.48%; 04: 10.7%; 05: 4.598%
| Z-Score (Absolute) |  | 01: 20.23; 02: 0.67; 03: 0.73; 04: 0.3; 05: 0
| Analytical results for Fe in OREAS 501b (Certified Value 4.08 wt.%) | [24 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 2
3
| 1 | 1 | 4.5103 | 4.2 | 4.28 | 3.8 | 3.91 | 3.75 | 3.95 | 4.44 | 2.73 | 4.17 | 3.77 | 4.04 | 4.05 | 4.0583 | 3.81 | 4.09 | 3.39 | 3.9 | 4.46 | 3.98 | 4.3 | 3.66 | 4.17 | 1
| 1 | 2 | 4.4891 | 4.21 | 4.27 | 3.87 | 3.87 | 3.84 | 4.16 | 4.6 | 2.72 | 4.27 | 3.87 | 4.03 | 4.03 | 4.1261 | 3.86 | 3.96 | 3.43 | 4.25 | 4.59 | 3.99 | 4.22 | 3.31 | 4.13 | #N/A
| 1 | 3 | 4.3927 | 4.22 | 4.39 | 3.79 | 3.88 | 3.79 | 4.11 | 4.37 | 2.7 | 4.17 | 3.8 | 4.05 | 4.02 | 4.0524 | 3.91 | 4.02 | 3.75 | 4.06 | 4.48 | 4.15 | 4.29 | 3.46 | 4.15 | 16
| 1 | 4 | 4.4572 | 4.11 | 4.48 | 3.81 | 3.86 | 3.73 | 4.23 | 4.51 | 2.75 | 4.2 | 3.87 | 4.13 | 3.9 | 4.1425 | 3.89 | 4.12 | 4.43 | 4.01 | 4.36 | 4.25 | 4.17 | 3.75 | 4.18 | 4.078
| 1 | 5 | 4.3509 | 4.02 | 4.34 | 3.83 | 3.84 | 3.79 | 4.09 | 4.5 | 2.71 | 4.27 | 3.76 | 4.15 | 3.89 | 4.1193 | 3.94 | 3.96 | 3.82 | 4.03 | 4.32 | 4.11 | 4.31 | 4.26 | 4.21 | 78
| 1 | 6 | 4.4397 | 4.21 | 4.31 | 3.91 | 3.84 | 3.75 | 4.06 | 4.39 | 2.63 | 4.24 | 3.87 | 4.15 | 3.9 | 4.075 | 3.99 | 3.99 | 3.58 | 4.06 | 4.54 | 4.36 | 4.17 | 4.25 | 4.33
| Mean |  | 4.44 | 4.1617 | 4.345 | 3.835 | 3.8667 | 3.775 | 4.1 | 4.4683 | 2.7067 | 4.22 | 3.8233 | 4.0917 | 3.965 | 4.0956 | 3.9 | 4.0233 | 3.7333 | 4.0517 | 4.4583 | 4.14 | 4.2433 | 3.7817 | 4.195
| Median |  | 4.4485 | 4.205 | 4.325 | 3.82 | 3.865 | 3.77 | 4.1 | 4.47 | 2.715 | 4.22 | 3.835 | 4.09 | 3.96 | 4.0971 | 3.9 | 4.005 | 3.665 | 4.045 | 4.47 | 4.13 | 4.255 | 3.705 | 4.175
| Std Dev. |  | 0.059708 | 0.080353 | 0.079183 | 0.046368 | 0.026583 | 0.039875 | 0.094657 | 0.085654 | 0.041312 | 0.046476 | 0.052789 | 0.057417 | 0.075565 | 0.038411 | 0.06261 | 0.067725 | 0.38119 | 0.11374 | 0.10323 | 0.14805 | 0.065013 | 0.39746 | 0.071484
| Rel.Std.Dev. |  | 1.345% | 1.931% | 1.822% | 1.209% | 0.6875% | 1.056% | 2.309% | 1.917% | 1.526% | 1.101% | 1.381% | 1.403% | 1.906% | 0.9379% | 1.605% | 1.683% | 10.21% | 2.807% | 2.315% | 3.576% | 1.532% | 10.51% | 1.704%
| PDM3 |  | 8.877% | 2.053% | 6.548% | -5.958% | -5.181% | -7.429% | 0.5403% | 9.573% | -33.63% | 3.483% | -6.244% | 0.336% | -2.77% | 0.4323% | -4.364% | -1.34% | -8.451% | -0.6449% | 9.327% | 1.521% | 4.055% | -7.266% | 2.87%
| Z-Score (Absolute) |  | 1.23 | 0.25 | 0.89 | 0.9 | 0.79 | 1.11 | 0.03 | 1.33 | 4.87 | 0.45 | 0.94 | 0 | 0.45 | 0.01 | 0.67 | 0.24 | 1.26 | 0.14 | 1.29 | 0.17 | 0.53 | 1.09 | 0.36
| Analytical results for Ga in OREAS 501b (Certified Value 8.7 ppm) | [24 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  | 2
3
| 1 | 1 | 8.2589 | 8.43 | 10 | 8.44 | 8.25 | 8 | 8.2 | 8.35 | 5 | 26 | 8 | 10 | 8.77 | 7.63 | 8.72 | < 10 | 8.6 | 8.5 | 9.5 | 27 | 13 |  |  | 1
| 1 | 2 | 8.6976 | 8.55 | 10 | 8.55 | 8.35 | 8 | 8.7 | 8.59 | 6 | 28 | 7.73 | 10 | 8.87 | 7.66 | 8.76 | < 10 | 9.07 | 8.5 | 9.5 | 28 | 14 |  |  | #N/A
| 1 | 3 | 8.4827 | 8.65 | 10.4 | 8.27 | 8.6 | 8 | 8.6 | 8.44 | 6 | 24 | 7.89 | 10 | 9.16 | 7.89 | 8.7 | < 10 | 8.65 | 8.5 | 9.4 | 24 | 13 |  |  | 16
| 1 | 4 | 8.2469 | 8.86 | 10.4 | 8.42 | 8.38 | 8 | 8.5 | 8.44 | 6 | 20 | 7.87 | 10 | 9.08 | 7.87 | 9.05 | 11 | 8.93 | 8.6 | 9.4 | 25 | 14 |  |  | 8.6957
| 1 | 5 | 8.2442 | 8.41 | 9.2 | 8.38 | 8.55 | 8 | 8.5 | 8.63 | 6 | 21 | 7.89 | < 10 | 9.08 | 8.63 | 8.65 | < 10 | 8.94 | 8.6 | 9.6 | 27 | 13 |  |  | 79
| 1 | 6 | 8.5415 | 8.3 | 9.4 | 8.83 | 8.43 | 8 | 8.5 | 8.33 | 6 | 23 | 7.99 | 10 | 8.84 | 8.09 | 8.04 | < 10 | 9.06 | 8.4 | 9.2 | 24 | 14
| Mean |  | 8.412 | 8.5333 | 9.9 | 8.4817 | 8.4267 | 8 | 8.5 | 8.4633 | 5.8333 | 23.6667 | 7.895 | 10 | 8.9667 | 7.9617 | 8.6533 | 11 | 8.875 | 8.5167 | 9.4333 | 25.8333 | 13.5
| Median |  | 8.3708 | 8.49 | 10 | 8.43 | 8.405 | 8 | 8.5 | 8.44 | 6 | 23.5 | 7.89 | 10 | 8.975 | 7.88 | 8.71 | 11 | 8.935 | 8.5 | 9.45 | 26 | 13.5
| Std Dev. |  | 0.19089 | 0.20047 | 0.502 | 0.19323 | 0.13003 | 0 | 0.16733 | 0.12291 | 0.40825 | 3.0111 | 0.097929 | 0 | 0.15946 | 0.36815 | 0.33212 |  | 0.20285 | 0.075277 | 0.13663 | 1.7224 | 0.54772
| Rel.Std.Dev. |  | 2.269% | 2.349% | 5.071% | 2.278% | 1.543% | 0% | 1.969% | 1.452% | 6.999% | 12.72% | 1.24% | 0% | 1.778% | 4.624% | 3.838% |  | 2.286% | 0.8839% | 1.448% | 6.667% | 4.057%
| PDM3 |  | -3.263% | -1.868% | 13.85% | -2.462% | -3.094% | -8.001% | -2.251% | -2.673% | -32.92% | 172.2% | -9.208% | 15% | 3.116% | -8.442% | -0.4875% | 26.5% | 2.062% | -2.059% | 8.482% | 197.1% | 55.25%
| Z-Score (Absolute) |  | 0.15 | 0 | 1.73 | 0.07 | 0.13 | 0.67 | 0.04 | 0.09 | 3.41 | 19.13 | 0.81 | 1.85 | 0.55 | 0.72 | 0.15 | 2.07 | 0.43 | 0.02 | 1.14 | 21.87 | 6.28
| Analytical results for Gd in OREAS 501b (Indicative Value 3.89 ppm) | [24 empty] | Indicative
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS |  |  | 2
2
| 1 | 1 | 01: < 5; 02: 3.76; 03: 3.6; 04: 3.86; 05: 3.9 |  |  | 1
| 1 | 2 | 01: < 5; 02: 2.63; 03: 3.56; 04: 4.24; 05: 3.9 |  |  | 7
| 1 | 3 | 01: < 5; 02: 3.95; 03: 3.73; 04: 4.13; 05: 4 |  |  | 16
| 1 | 4 | 01: < 5; 02: 3.76; 03: 3.75; 04: 4.17; 05: 4 |  |  | 3.8884
| 1 | 5 | 01: < 5; 02: 3.92; 03: 3.58; 04: 4.17; 05: 4 |  |  | 13
| 1 | 6 | 01: < 5; 02: 3.82; 03: 3.68; 04: 4; 05: 4
| Mean |  | 02: 3.64; 03: 3.65; 04: 4.095; 05: 3.9667
| Median |  | 02: 3.79; 03: 3.64; 04: 4.15; 05: 4
| Std Dev. |  | 02: 0.50116; 03: 0.080994; 04: 0.13982; 05: 0.05164
| Rel.Std.Dev. |  | 02: 13.77%; 03: 2.219%; 04: 3.414%; 05: 1.302%
| PDM3 |  | 02: -6.389%; 03: -6.131%; 04: 5.313%; 05: 2.012%
| Z-Score (Absolute) |  | 01: 2.2; 02: 0.7; 03: 0.67; 04: 0.27; 05: 0
| Analytical results for Ge in OREAS 501b (Certified Value 0.2 ppm) | [24 empty] | Certified
Ge | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 10: Lab; 11: Lab; 12: Lab |  |  | 1
| No | No | 01: 04; 02: 05; 03: 07; 04: 12; 05: 16; 06: 17; 07: 18; 08: 19; 10: 20; 11: 23; 12: 25 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*MS; 07: AR*MS; 08: AR*OES; 10: AR*MS; 11: AR*OES/MS; 12: AR*OES/MS |  |  | 2
3
| 1 | 1 | 01: 0.19; 02: 0.2; 03: 0.2; 04: 0.23; 05: 0.25326; 06: 0.18; 07: 0.21; 08: 3; 10: 0.15; 11: < 0.5; 12: < 0.1 |  |  | 1
| 1 | 2 | 01: 0.19; 02: 0.19; 03: 0.2; 04: 0.2; 05: 0.24793; 06: 0.19; 07: 0.23; 08: 4; 10: 0.16; 11: < 0.5; 12: < 0.1 |  |  | 5
| 1 | 3 | 01: 0.18; 02: 0.21; 03: 0.2; 04: 0.21; 05: 0.24479; 06: 0.2; 07: 0.21; 08: < 1; 10: 0.14; 11: < 0.5; 12: < 0.1 |  |  | 16
| 1 | 4 | 01: 0.18; 02: 0.18; 03: 0.2; 04: 0.21; 05: 0.24023; 06: 0.2; 07: 0.23; 08: 4; 10: 0.14; 11: < 0.5; 12: < 0.1 |  |  | 0.19994
| 1 | 5 | 01: 0.19; 02: 0.18; 03: 0.2; 04: 0.21; 05: 0.25311; 06: 0.18; 07: 0.23; 08: < 1; 10: 0.14; 11: < 0.5; 12: < 0.1 |  |  | 80
| 1 | 6 | 01: 0.2; 02: 0.19; 03: 0.2; 04: 0.21; 05: 0.26389; 06: 0.2; 07: 0.21; 08: < 1; 10: 0.15; 11: < 0.5; 12: < 0.1
| Mean |  | 01: 0.18833; 02: 0.19167; 03: 0.2; 04: 0.21167; 05: 0.25054; 06: 0.19167; 07: 0.22; 08: 3.6667; 10: 0.14667
| Median |  | 01: 0.19; 02: 0.19; 03: 0.2; 04: 0.21; 05: 0.25052; 06: 0.195; 07: 0.22; 08: 4; 10: 0.145
| Std Dev. |  | 01: 0.0075277; 02: 0.01169; 03: 3.0405e-17; 04: 0.0098319; 05: 0.0082255; 06: 0.0098319; 07: 0.010954; 08: 0.57735; 10: 0.008165
| Rel.Std.Dev. |  | 01: 3.997%; 02: 6.099%; 03: 1.52e-14%; 04: 4.645%; 05: 3.283%; 06: 5.13%; 07: 4.979%; 08: 15.75%; 10: 5.567%
| PDM3 |  | 01: -5.807%; 02: -4.14%; 03: 0.02779%; 04: 5.863%; 05: 25.3%; 06: -4.14%; 07: 10.03%; 08: 1734%; 10: -26.65%
| Z-Score (Absolute) |  | 01: 0.39; 02: 0.28; 03: 0; 04: 0.39; 05: 1.7; 06: 0.28; 07: 0.67; 08: 71.93; 10: 1.8; 11: 10.11; 12: 3.37
| Analytical results for Hf in OREAS 501b (Certified Value 0.43 ppm) | [24 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | [9 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [9 empty] | 2
2
| 1 | 1 | 0.07 | 0.2 | 0.42 | 0.43 | 0.42 | 0.49 | 0.41 | 0.57391 | 0.47 | 0.46 | < 1 | 0.36 | 0.4 | 0.28 | [9 empty] | 1
| 1 | 2 | 0.04 | 0.2 | 0.42 | 0.48 | 0.43 | 0.5 | 0.41 | 0.53216 | 0.47 | 0.46 | < 1 | 0.36 | 0.4 | 0.31 | [9 empty] | #N/A
| 1 | 3 | 0.06 | 0.2 | 0.39 | 0.47 | 0.45 | 0.48 | 0.38 | 0.51771 | 0.45 | 0.45 | < 1 | 0.36 | 0.4 | 0.3 | [9 empty] | 16
| 1 | 4 | 0.03 | 0.2 | 0.43 | 0.43 | 0.47 | 0.63 | 0.4 | 0.52051 | 0.47 | 0.47 | 2 | 0.34 | 0.4 | 0.31 | [9 empty] | 0.42758
| 1 | 5 | 0.03 | 0.2 | 0.41 | 0.43 | 0.43 | 0.53 | 0.41 | 0.51245 | 0.46 | 0.45 | < 1 | 0.35 | 0.4 | 0.31 | [9 empty] | 81
| 1 | 6 | 0.03 | 0.2 | 0.43 | 0.44 | 0.46 | 0.5 | 0.43 | 0.49252 | 0.47 | 0.41 | < 1 | 0.36 | 0.4 | 0.32
| Mean |  | 0.043333 | 0.2 | 0.41667 | 0.44667 | 0.44333 | 0.52167 | 0.40667 | 0.52488 | 0.465 | 0.45 | 2 | 0.355 | 0.4 | 0.305
| Median |  | 0.035 | 0.2 | 0.42 | 0.435 | 0.44 | 0.5 | 0.41 | 0.51911 | 0.47 | 0.455 | 2 | 0.36 | 0.4 | 0.31
| Std Dev. |  | 0.017512 | 3.0405e-17 | 0.015055 | 0.022509 | 0.019664 | 0.055648 | 0.01633 | 0.027311 | 0.0083666 | 0.020976 |  | 0.0083666 | 6.0809e-17 | 0.013784
| Rel.Std.Dev. |  | 40.41% | 1.52e-14% | 3.613% | 5.039% | 4.435% | 10.67% | 4.016% | 5.203% | 1.799% | 4.661% |  | 2.357% | 1.52e-14% | 4.519%
| PDM3 |  | -89.87% | -53.23% | -2.553% | 4.463% | 3.684% | 22% | -4.892% | 22.75% | 8.751% | 5.243% | 367.7% | -16.97% | -6.451% | -28.67%
| Z-Score (Absolute) |  | 4.74 | 2.82 | 0.16 | 0.2 | 0.16 | 1.12 | 0.29 | 1.16 | 0.43 | 0.25 | 9.03 | 0.92 | 0.37 | 1.53
| Analytical results for Hg in OREAS 501b (Indicative Value 0.019 ppm) | [24 empty] | Indicative
Hg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 24 | 25 | 27 |  |  |  |  | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 3
3
| 1 | 1 | < 0.01 | 0.03 | 0.02 | 0.01 | 0.02 | 0.03 | 0.01 | < 5 | < 1 | 0.02 | 1 | 0.018 | 0.02 | 0.02 | < 1 | 0.03 | 0.07 | 0.01 | < 5 |  |  |  |  | 1
| 1 | 2 | < 0.01 | 0.03 | 0.02 | 0.03 | 0.02 | 0.03 | 0.01 | < 5 | < 1 | 0.01 | 1 | 0.009 | 0.02 | 0.02 | < 1 | 0.02 | 0.03 | < 0.01 | < 5 |  |  |  |  | 3
| 1 | 3 | < 0.01 | 0.02 | 0.02 | 0.02 | 0.01 | 0.03 | 0.03 | < 5 | < 1 | < 0.01 | 1 | 0.022 | 0.02 | 0.02 | 1 | 0.02 | 0.05 | 0.01 | < 5 |  |  |  |  | 16
| 1 | 4 | 0.1 | 0.02 | 0.02 | 0.01 | 0.02 | 0.03 | 0.02 | < 5 | < 1 | < 0.01 | 1 | 0.015 | 0.02 | 0.02 | < 1 | 0.01 | 0.03 | < 0.01 | < 5 |  |  |  |  | 0.018778
| 1 | 5 | 0.08 | 0.04 | 0.02 | 0.01 | 0.02 | 0.04 | 0.03 | < 5 | < 1 | 0.02 | 1 | 0.011 | 0.02 | 0.02 | < 1 | 0.01 | 0.09 | < 0.01 | < 5 |  |  |  |  | 14
| 1 | 6 | 0.05 | 0.02 | 0.02 | 0.02 | 0.01 | 0.03 | < 0.01 | < 5 | < 1 | 0.01 | < 1 | 0.013 | 0.02 | 0.02 | < 1 | 0.02 | 0.03 | < 0.01 | < 5
| Mean |  | 0.076667 | 0.026667 | 0.02 | 0.016667 | 0.016667 | 0.031667 | 0.02 |  |  | 0.015 | 1 | 0.014667 | 0.02 | 0.02 | 1 | 0.018333 | 0.05 | 0.01
| Median |  | 0.08 | 0.025 | 0.02 | 0.015 | 0.02 | 0.03 | 0.02 |  |  | 0.015 | 1 | 0.014 | 0.02 | 0.02 | 1 | 0.02 | 0.04 | 0.01
| Std Dev. |  | 0.025166 | 0.008165 | 0 | 0.008165 | 0.005164 | 0.0040825 | 0.01 |  |  | 0.0057735 | 0 | 0.004761 | 0 | 0 |  | 0.0075277 | 0.025298 | 0
| Rel.Std.Dev. |  | 32.83% | 30.62% | 0% | 48.99% | 30.98% | 12.89% | 50% |  |  | 38.49% | 0% | 32.46% | 0% | 0% |  | 41.06% | 50.6% | 0%
| PDM3 |  | 308.3% | 42.01% | 6.509% | -11.24% | -11.24% | 68.64% | 6.509% |  |  | -20.12% | 5225% | -21.89% | 6.509% | 6.509% | 5225% | -2.367% | 166.3% | -46.75%
| Z-Score (Absolute) |  | 2.36 | 0.67 | 0 | 0.34 | 0.34 | 1.18 | 0.17 | 503.71 | 99.12 | 0.67 | 99.12 | 0.54 | 0 | 0 | 99.12 | 0.17 | 3.03 | 1.01 | 503.71
| Analytical results for Ho in OREAS 501b (Indicative Value 0.61 ppm) | [24 empty] | Indicative
Ho | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  |  |  |  | 1
| No | No | 02: 01; 03: 02; 04: 03; 05: 08; 06: 23 |  |  |  |  | ppm
|  |  | 02: AR*MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*MS; 06: AR*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 02: < 5; 03: 0.62; 04: 0.565; 05: 0.63; 06: 0.6 |  |  |  |  | 1
| 1 | 2 | 02: < 5; 03: 0.62; 04: 0.555; 05: 0.67; 06: 0.6 |  |  |  |  | 9
| 1 | 3 | 02: < 5; 03: 0.65; 04: 0.58; 05: 0.64; 06: 0.6 |  |  |  |  | 16
| 1 | 4 | 02: < 5; 03: 0.63; 04: 0.585; 05: 0.66; 06: 0.6 |  |  |  |  | 0.61229
| 1 | 5 | 02: < 5; 03: 0.63; 04: 0.575; 05: 0.66; 06: 0.6 |  |  |  |  | 15
| 1 | 6 | 02: < 5; 03: 0.63; 04: 0.555; 05: 0.64; 06: 0.6
| Mean |  | 03: 0.63; 04: 0.56917; 05: 0.65; 06: 0.6
| Median |  | 03: 0.63; 04: 0.57; 05: 0.65; 06: 0.6
| Std Dev. |  | 03: 0.010954; 04: 0.012813; 05: 0.015492; 06: 0
| Rel.Std.Dev. |  | 03: 1.739%; 04: 2.251%; 05: 2.383%; 06: 0%
| PDM3 |  | 03: 2.892%; 04: -7.043%; 05: 6.159%; 06: -2.007%
| Z-Score (Absolute) |  | 02: 98.22; 03: 0; 04: 1.37; 05: 0.45; 06: 0.67
| Analytical results for In in OREAS 501b (Certified Value 0.19 ppm) | [24 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 27 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | [7 empty] | 2
3
| 1 | 1 | < 0.5 | 0.177 | 0.175 | 0.189 | 0.198 | 0.18 | 0.185 | 0.192 | 0.189 | 0.221 | 0.197 | < 1 | 0.184 | 0.2 | 0.17 | < 5 | [7 empty] | 1
| 1 | 2 | < 0.5 | 0.174 | 0.18 | 0.192 | 0.198 | 0.19 | 0.195 | 0.181 | 0.194 | 0.214 | 0.203 | < 1 | 0.199 | 0.2 | 0.15 | < 5 | [7 empty] | #N/A
| 1 | 3 | < 0.5 | 0.177 | 0.18 | 0.186 | 0.196 | 0.19 | 0.207 | 0.191 | 0.21 | 0.217 | 0.187 | < 1 | 0.192 | 0.2 | 0.16 | < 5 | [7 empty] | 16
| 1 | 4 | < 0.5 | 0.182 | 0.18 | 0.186 | 0.2 | 0.19 | 0.192 | 0.187 | 0.205 | 0.217 | 0.193 | 3 | 0.187 | 0.2 | 0.15 | < 5 | [7 empty] | 0.19344
| 1 | 5 | < 0.5 | 0.176 | 0.175 | 0.191 | 0.204 | 0.19 | 0.215 | 0.188 | 0.207 | 0.214 | 0.199 | < 1 | 0.181 | 0.2 | 0.15 | < 5 | [7 empty] | 82
| 1 | 6 | < 0.5 | 0.178 | 0.17 | 0.19 | 0.206 | 0.19 | 0.213 | 0.194 | 0.198 | 0.216 | 0.179 | < 1 | 0.185 | 0.2 | 0.15 | < 5
| Mean |  |  | 0.17733 | 0.17667 | 0.189 | 0.20033 | 0.18833 | 0.20117 | 0.18883 | 0.2005 | 0.2165 | 0.193 | 3 | 0.188 | 0.2 | 0.155
| Median |  |  | 0.177 | 0.1775 | 0.1895 | 0.199 | 0.19 | 0.201 | 0.1895 | 0.2015 | 0.2165 | 0.195 | 3 | 0.186 | 0.2 | 0.15
| Std Dev. |  |  | 0.0026583 | 0.0040825 | 0.0025298 | 0.0038816 | 0.0040825 | 0.012238 | 0.0046224 | 0.008167 | 0.0025884 | 0.0087636 |  | 0.0065115 | 3.0405e-17 | 0.0083666
| Rel.Std.Dev. |  |  | 1.499% | 2.311% | 1.339% | 1.938% | 2.168% | 6.083% | 2.448% | 4.073% | 1.196% | 4.541% |  | 3.464% | 1.52e-14% | 5.398%
| PDM3 |  |  | -8.329% | -8.673% | -2.298% | 3.561% | -2.642% | 3.992% | -2.384% | 3.647% | 11.92% | -0.2298% | 1451% | -2.814% | 3.389% | -19.87%
| Z-Score (Absolute) |  | 24.56 | 1.55 | 1.6 | 0.61 | 0.31 | 0.66 | 0.38 | 0.62 | 0.32 | 1.62 | 0.28 | 91.99 | 0.69 | 0.28 | 3.36 | 388.69
| Analytical results for K in OREAS 501b (Certified Value 1.08 wt.%) | [24 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 2
3
| 1 | 1 | 1.3315 | 1.107 | 1.05 | 1.02 | 1.06 | 1.06 | 1.09 | 1.1051 | 0.24 | 0.98 | 0.92 | 1.05 | 0.95 | 1.0937 | 1.03 | 1.11 | 0.91 | 1.1 | 1.27 | 1.08 | 1.25 | 1.07 | 1.16 | 1
| 1 | 2 | 1.3785 | 1.113 | 1.04 | 1.05 | 1.05 | 1.1 | 1.11 | 1.1362 | 0.23 | 1 | 0.92 | 1.04 | 0.98 | 1.1021 | 1.05 | 1.07 | 0.93 | 1.09 | 1.31 | 1.09 | 1.22 | 0.97 | 1.14 | #N/A
| 1 | 3 | 1.2918 | 1.1 | 1.06 | 1.02 | 1.07 | 1.08 | 1.1 | 1.107 | 0.24 | 1.01 | 0.93 | 1.05 | 0.98 | 1.1024 | 1.06 | 1.11 | 1.03 | 1.1 | 1.3 | 1.13 | 1.23 | 1.16 | 1.15 | 16
| 1 | 4 | 1.3082 | 1.14 | 1.08 | 1.04 | 1.07 | 1.07 | 1.11 | 1.1152 | 0.23 | 0.99 | 0.94 | 1.07 | 0.94 | 1.1189 | 1.05 | 1.12 | 1.18 | 1.09 | 1.25 | 1.15 | 1.2 | 1.28 | 1.17 | 1.0848
| 1 | 5 | 1.2645 | 1.128 | 1.1 | 1.04 | 1.05 | 1.09 | 1.08 | 1.1202 | 0.24 | 1 | 0.93 | 1.08 | 0.94 | 1.1212 | 1.07 | 1.09 | 1.04 | 1.09 | 1.23 | 1.11 | 1.25 | 1.19 | 1.16 | 83
| 1 | 6 | 1.335 | 1.11 | 1.09 | 1.06 | 1.05 | 1.08 | 1.1 | 1.0943 | 0.23 | 1 | 0.95 | 1.08 | 0.95 | 1.0838 | 1.08 | 1.08 | 0.97 | 1.1 | 1.29 | 1.18 | 1.21 | 1.21 | 1.2
| Mean |  | 1.3182 | 1.1163 | 1.07 | 1.0383 | 1.0583 | 1.08 | 1.0983 | 1.113 | 0.235 | 0.99667 | 0.93167 | 1.0617 | 0.95667 | 1.1037 | 1.0567 | 1.0967 | 1.01 | 1.095 | 1.275 | 1.1233 | 1.2267 | 1.1467 | 1.1633
| Median |  | 1.3198 | 1.1115 | 1.07 | 1.04 | 1.055 | 1.08 | 1.1 | 1.1111 | 0.235 | 1 | 0.93 | 1.06 | 0.95 | 1.1022 | 1.055 | 1.1 | 1 | 1.095 | 1.28 | 1.12 | 1.225 | 1.175 | 1.16
| Std Dev. |  | 0.039436 | 0.014841 | 0.023664 | 0.016021 | 0.0098319 | 0.014142 | 0.01169 | 0.014443 | 0.0054772 | 0.010328 | 0.01169 | 0.017224 | 0.018619 | 0.014418 | 0.017512 | 0.019664 | 0.098184 | 0.0054772 | 0.030822 | 0.037771 | 0.020656 | 0.11039 | 0.020656
| Rel.Std.Dev. |  | 2.992% | 1.329% | 2.212% | 1.543% | 0.929% | 1.309% | 1.064% | 1.298% | 2.331% | 1.036% | 1.255% | 1.622% | 1.946% | 1.306% | 1.657% | 1.793% | 9.721% | 0.5002% | 2.417% | 3.362% | 1.684% | 9.627% | 1.776%
| PDM3 |  | 21.52% | 2.906% | -1.365% | -4.284% | -2.44% | -0.4431% | 1.247% | 2.599% | -78.34% | -8.125% | -14.12% | -2.133% | -11.81% | 1.739% | -2.594% | 1.093% | -6.896% | 0.9396% | 17.53% | 3.551% | 13.08% | 5.702% | 7.239%
| Z-Score (Absolute) |  | 3.93 | 0.38 | 0.44 | 1 | 0.64 | 0.26 | 0.06 | 0.32 | 15.13 | 1.73 | 2.87 | 0.59 | 2.43 | 0.15 | 0.67 | 0.03 | 1.5 | 0 | 3.17 | 0.5 | 2.32 | 0.91 | 1.2
| Analytical results for La in OREAS 501b (Certified Value 29.3 ppm) | [24 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
2
| 1 | 1 | 31.776 | 30.09 | 29.9 | 29.6 | 29.4 | 29 | 30 | 29.75 | 8.5 | 25 | 29 | 30 | 30 | 27.49 | 33.4 | 32.2 | 23 | 27.9 | 30 | 26.4 | 27 | 27 |  | 1
| 1 | 2 | 32.278 | 29.96 | 29.9 | 29 | 32 | 30 | 29.2 | 32.52 | 9 | 26 | 29 | 30 | 30 | 27.02 | 34.4 | 32.5 | 23 | 29.7 | 32 | 25.5 | 25 | 28 |  | #N/A
| 1 | 3 | 31.471 | 30.35 | 30.9 | 27.9 | 31.6 | 29 | 28.3 | 30.92 | 8.6 | 26 | 31 | 30 | 30 | 28.05 | 34.9 | 31.6 | 25 | 30.4 | 30 | 25.7 | 27 | 28 |  | 16
| 1 | 4 | 30.857 | 29.83 | 31.4 | 28.5 | 32.1 | 29 | 28 | 31.73 | 8.8 | 26 | 30 | 30 | 30 | 29.29 | 35 | 31.9 | 32 | 31 | 29 | 24.6 | 29 | 29 |  | 29.3216
| 1 | 5 | 30.717 | 30.4 | 29.3 | 27.6 | 30.8 | 30 | 28.1 | 32.26 | 8.9 | 26 | 30 | 30 | 20 | 27.35 | 30.5 | 33.1 | 24 | 28.3 | 28 | 26.2 | 28 | 29 |  | 84
| 1 | 6 | 31.931 | 30.22 | 29.2 | 28.9 | 30 | 28 | 28.8 | 30.55 | 8.8 | 26 | 29 | 30 | 30 | 26.83 | 34.5 | 29.7 | 22 | 30.4 | 32 | 26.9 | 29 | 29
| Mean |  | 31.505 | 30.1417 | 30.1 | 28.5833 | 30.9833 | 29.1667 | 28.7333 | 31.2883 | 8.7667 | 25.8333 | 29.6667 | 30 | 28.3333 | 27.6717 | 33.7833 | 31.8333 | 24.8333 | 29.6167 | 30.1667 | 25.8833 | 27.5 | 28.3333
| Median |  | 31.6235 | 30.155 | 29.9 | 28.7 | 31.2 | 29 | 28.55 | 31.325 | 8.8 | 26 | 29.5 | 30 | 30 | 27.42 | 34.45 | 32.05 | 23.5 | 30.05 | 30 | 25.95 | 27.5 | 28.5
| Std Dev. |  | 0.61554 | 0.2232 | 0.87864 | 0.7414 | 1.1107 | 0.75277 | 0.76855 | 1.068 | 0.18619 | 0.40825 | 0.8165 | 0 | 4.0825 | 0.89796 | 1.7058 | 1.1656 | 3.656 | 1.2513 | 1.6021 | 0.80353 | 1.5166 | 0.8165
| Rel.Std.Dev. |  | 1.954% | 0.7405% | 2.919% | 2.594% | 3.585% | 2.581% | 2.675% | 3.414% | 2.124% | 1.58% | 2.752% | 0% | 14.41% | 3.245% | 5.049% | 3.662% | 14.72% | 4.225% | 5.311% | 3.104% | 5.515% | 2.882%
| PDM3 |  | 7.447% | 2.797% | 2.655% | -2.518% | 5.667% | -0.5283% | -2.006% | 6.708% | -70.1% | -11.9% | 1.177% | 2.314% | -3.37% | -5.627% | 15.22% | 8.566% | -15.31% | 1.006% | 2.882% | -11.73% | -6.212% | -3.37%
| Z-Score (Absolute) |  | 1.08 | 0.38 | 0.36 | 0.41 | 0.81 | 0.11 | 0.34 | 0.97 | 10.5 | 1.81 | 0.14 | 0.31 | 0.54 | 0.88 | 2.23 | 1.24 | 2.32 | 0.11 | 0.39 | 1.79 | 0.96 | 0.54
| Analytical results for Li in OREAS 501b (Certified Value 31.7 ppm) | [24 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
2
| 1 | 1 | 34.364 | 32 | 34.4 | 30 | 32.8 | 27 | 33.1 | 15 | 28 | 33.7 | 31.0217 | 29.4 | 33.7 | 28 | 28.9 | 30.4 | 30 | 32 | 27 | 21 |  |  |  | 1
| 1 | 2 | 35.339 | 32.2 | 33 | 31.2 | 35.1 | 28 | 33.8 | 15 | 29 | 31.7 | 31.6878 | 29.8 | 34.3 | 29 | 30.1 | 30.8 | 30 | 31 | 21 | 21 |  |  |  | #N/A
| 1 | 3 | 33.673 | 32.6 | 35.1 | 29.9 | 35.7 | 28 | 33.4 | 15 | 29 | 33 | 31.9167 | 30.6 | 33.5 | 32 | 30.3 | 31.4 | 31 | 33 | 32 | 21 |  |  |  | 16
| 1 | 4 | 33.886 | 32.3 | 33.5 | 30.6 | 34.3 | 28 | 33.5 | 15 | 28 | 31.6 | 32.2705 | 30.5 | 34.3 | 38 | 31.2 | 32 | 32 | 32 | 33 | 21 |  |  |  | 31.7016
| 1 | 5 | 33.748 | 32.8 | 34.6 | 30.3 | 35.4 | 27 | 34.4 | 15 | 29 | 32.2 | 31.277 | 37.1 | 33.2 | 32 | 31.2 | 30.8 | 31 | 32 | 33 | 22 |  |  |  | 85
| 1 | 6 | 34.6 | 32.4 | 33.6 | 30.4 | 34.5 | 28 | 33 | 15 | 29 | 32.5 | 31.2297 | 31.5 | 31.3 | 30 | 30.9 | 30.2 | 32 | 32 | 30 | 23
| Mean |  | 34.2683 | 32.3833 | 34.0333 | 30.4 | 34.6333 | 27.6667 | 33.5333 | 15 | 28.6667 | 32.45 | 31.5672 | 31.4833 | 33.3833 | 31.5 | 30.4333 | 30.9333 | 31 | 32 | 29.3333 | 21.5
| Median |  | 34.125 | 32.35 | 34 | 30.35 | 34.8 | 28 | 33.45 | 15 | 29 | 32.35 | 31.4824 | 30.55 | 33.6 | 31 | 30.6 | 30.8 | 31 | 32 | 31 | 21
| Std Dev. |  | 0.63825 | 0.28577 | 0.79162 | 0.46904 | 1.0424 | 0.5164 | 0.51251 | 0 | 0.5164 | 0.80187 | 0.47476 | 2.8449 | 1.1107 | 3.5637 | 0.88015 | 0.66533 | 0.89443 | 0.63246 | 4.6762 | 0.83666
| Rel.Std.Dev. |  | 1.863% | 0.8825% | 2.326% | 1.543% | 3.01% | 1.866% | 1.528% | 0% | 1.801% | 2.471% | 1.504% | 9.036% | 3.327% | 11.31% | 2.892% | 2.151% | 2.885% | 1.976% | 15.94% | 3.891%
| PDM3 |  | 8.097% | 2.15% | 7.355% | -4.106% | 9.248% | -12.73% | 5.778% | -52.68% | -9.573% | 2.361% | -0.4239% | -0.6885% | 5.305% | -0.6359% | -4.001% | -2.423% | -2.213% | 0.9413% | -7.471% | -32.18%
| Z-Score (Absolute) |  | 1.26 | 0.4 | 1.15 | 0.49 | 1.42 | 1.73 | 0.92 | 7.46 | 1.28 | 0.43 | 0.03 | 0 | 0.86 | 0 | 0.48 | 0.25 | 0.22 | 0.23 | 0.98 | 4.52
| Analytical results for Lu in OREAS 501b (Certified Value 0.21 ppm) | [24 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 07: Lab; 08: Lab; 09: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 07; 05: 08; 07: 16; 08: 23; 09: 25 |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS; 07: AR*OES/MS; 08: AR*OES/MS; 09: AR*OES/MS |  |  |  | 2
3
| 1 | 1 | 01: < 5; 02: 0.221; 03: 0.19; 04: 0.21; 05: 0.218; 07: 0.23913; 08: 0.2; 09: 0.19 |  |  |  | 1
| 1 | 2 | 01: < 5; 02: 0.228; 03: 0.19; 04: 0.21; 05: 0.231; 07: 0.24066; 08: 0.2; 09: 0.2 |  |  |  | #N/A
| 1 | 3 | 01: < 5; 02: 0.226; 03: 0.195; 04: 0.22; 05: 0.218; 07: 0.24583; 08: 0.2; 09: 0.2 |  |  |  | 16
| 1 | 4 | 01: < 5; 02: 0.227; 03: 0.2; 04: 0.22; 05: 0.234; 07: 0.24316; 08: 0.21; 09: 0.19 |  |  |  | 0.21483
| 1 | 5 | 01: < 5; 02: 0.226; 03: 0.19; 04: 0.22; 05: 0.235; 07: 0.24378; 08: 0.2; 09: 0.2 |  |  |  | 86
| 1 | 6 | 01: < 5; 02: 0.225; 03: 0.185; 04: 0.23; 05: 0.225; 07: 0.23611; 08: 0.21; 09: 0.2
| Mean |  | 02: 0.2255; 03: 0.19167; 04: 0.21833; 05: 0.22683; 07: 0.24145; 08: 0.20333; 09: 0.19667
| Median |  | 02: 0.226; 03: 0.19; 04: 0.22; 05: 0.228; 07: 0.24191; 08: 0.2; 09: 0.2
| Std Dev. |  | 02: 0.002429; 03: 0.005164; 04: 0.0075277; 05: 0.007679; 07: 0.0035232; 08: 0.005164; 09: 0.005164
| Rel.Std.Dev. |  | 02: 1.077%; 03: 2.694%; 04: 3.448%; 05: 3.385%; 07: 1.459%; 08: 2.54%; 09: 2.626%
| PDM3 |  | 02: 4.969%; 03: -10.78%; 04: 1.633%; 05: 5.589%; 07: 12.39%; 08: -5.35%; 09: -8.453%
| Z-Score (Absolute) |  | 01: 169.07; 02: 0.13; 03: 1.07; 04: 0.13; 05: 0.17; 07: 0.69; 08: 0.66; 09: 0.89
| Analytical results for Mg in OREAS 501b (Certified Value 1.21 wt.%) | [24 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 2
3
| 1 | 1 | 1.2702 | 1.28 | 1.19 | 1.17 | 1.18 | 1.15 | 1.18 | 1.21 | 0.81 | 1.14 | 1.18 | 1.19 | 1.07 | 1.2342 | 1.15 | 1.23 | 1 | 1.14 | 1.24 | 1.2 | 1.25 | 1.15 | 1.09 | 1
| 1 | 2 | 1.2669 | 1.28 | 1.19 | 1.17 | 1.17 | 1.2 | 1.2 | 1.25 | 0.82 | 1.18 | 1.2 | 1.18 | 1.15 | 1.2555 | 1.17 | 1.18 | 1.02 | 1.25 | 1.28 | 1.2 | 1.24 | 1.06 | 1.08 | #N/A
| 1 | 3 | 1.2532 | 1.28 | 1.22 | 1.15 | 1.18 | 1.17 | 1.19 | 1.2 | 0.82 | 1.18 | 1.21 | 1.19 | 1.14 | 1.2223 | 1.19 | 1.2 | 1.11 | 1.2 | 1.25 | 1.24 | 1.22 | 1.27 | 1.08 | 16
| 1 | 4 | 1.2532 | 1.31 | 1.27 | 1.15 | 1.17 | 1.16 | 1.21 | 1.22 | 0.82 | 1.16 | 1.22 | 1.21 | 1.11 | 1.2666 | 1.17 | 1.24 | 1.33 | 1.18 | 1.22 | 1.26 | 1.2 | 1.29 | 1.09 | 1.2074
| 1 | 5 | 1.2288 | 1.3 | 1.23 | 1.15 | 1.16 | 1.17 | 1.18 | 1.22 | 0.82 | 1.18 | 1.2 | 1.21 | 1.11 | 1.2562 | 1.19 | 1.19 | 1.13 | 1.19 | 1.21 | 1.22 | 1.23 | 1.3 | 1.09 | 87
| 1 | 6 | 1.2434 | 1.28 | 1.22 | 1.18 | 1.17 | 1.16 | 1.21 | 1.19 | 0.81 | 1.17 | 1.22 | 1.21 | 1.11 | 1.238 | 1.21 | 1.19 | 1.06 | 1.19 | 1.26 | 1.29 | 1.19 | 1.32 | 1.13
| Mean |  | 1.2526 | 1.2883 | 1.22 | 1.1617 | 1.1717 | 1.1683 | 1.195 | 1.215 | 0.81667 | 1.1683 | 1.205 | 1.1983 | 1.115 | 1.2455 | 1.18 | 1.205 | 1.1083 | 1.1917 | 1.2433 | 1.235 | 1.2217 | 1.2317 | 1.0933
| Median |  | 1.2532 | 1.28 | 1.22 | 1.16 | 1.17 | 1.165 | 1.195 | 1.215 | 0.82 | 1.175 | 1.205 | 1.2 | 1.11 | 1.2468 | 1.18 | 1.195 | 1.085 | 1.19 | 1.245 | 1.23 | 1.225 | 1.28 | 1.09
| Std Dev. |  | 0.015271 | 0.013292 | 0.029665 | 0.013292 | 0.0075277 | 0.017224 | 0.013784 | 0.020736 | 0.005164 | 0.016021 | 0.015166 | 0.013292 | 0.028107 | 0.016622 | 0.020976 | 0.02429 | 0.11957 | 0.035449 | 0.02582 | 0.035637 | 0.023166 | 0.10342 | 0.018619
| Rel.Std.Dev. |  | 1.219% | 1.032% | 2.432% | 1.144% | 0.6425% | 1.474% | 1.153% | 1.707% | 0.6323% | 1.371% | 1.259% | 1.109% | 2.521% | 1.335% | 1.778% | 2.016% | 10.79% | 2.975% | 2.077% | 2.886% | 1.896% | 8.397% | 1.703%
| PDM3 |  | 3.747% | 6.706% | 1.046% | -3.785% | -2.957% | -3.233% | -1.025% | 0.6318% | -32.36% | -3.233% | -0.1964% | -0.7486% | -7.651% | 3.156% | -2.267% | -0.1964% | -8.203% | -1.301% | 2.979% | 2.288% | 1.184% | 2.012% | -9.445%
| Z-Score (Absolute) |  | 1.22 | 2.02 | 0.49 | 0.82 | 0.6 | 0.67 | 0.07 | 0.37 | 8.58 | 0.67 | 0.15 | 0 | 1.87 | 1.06 | 0.41 | 0.15 | 2.02 | 0.15 | 1.01 | 0.82 | 0.52 | 0.75 | 2.36
| Analytical results for Mn in OREAS 501b (Certified Value 393 ppm) | [24 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 0
0
| 1 | 1 | 427.87 | 417 | 406 | 364 | 380 | 384 | 359 | 395 | 659 | 382 | 382 | 392 | 387 | 401.1304 | 359 | 397 | 339 | 389 | 400.4 | 401 | 400 | 452 | 400 | 1
| 1 | 2 | 435.62 | 424 | 402 | 371 | 377 | 388 | 366 | 406 | 668 | 387 | 387 | 390 | 367 | 407.1162 | 362 | 383 | 345 | 390 | 399.7 | 394 | 400 | 368 | 400 | #N/A
| 1 | 3 | 419.94 | 422 | 414 | 359 | 381 | 391 | 363 | 391 | 664 | 389 | 389 | 392 | 365 | 398.2083 | 368 | 396 | 377 | 394 | 399.6 | 410 | 400 | 413 | 400 | 16
| 1 | 4 | 420.77 | 431 | 424 | 362 | 376 | 372 | 370 | 396 | 662 | 383 | 391 | 398 | 360 | 407.1484 | 365 | 397 | 455 | 394 | 399.2 | 425 | 400 | 435 | 420 | 392.9451
| 1 | 5 | 410.57 | 427 | 387 | 366 | 374 | 384 | 361 | 399 | 666 | 387 | 387 | 400 | 354 | 403.6618 | 373 | 387 | 385 | 389 | 403.4 | 406 | 400 | 443 | 420 | 88
| 1 | 6 | 424.11 | 422 | 391 | 373 | 376 | 388 | 369 | 388 | 660 | 383 | 393 | 402 | 354 | 400.1068 | 376 | 386 | 360 | 394 | 396.2 | 432 | 400 | 403 | 420
| Mean |  | 423.1467 | 423.8333 | 404 | 365.8333 | 377.3333 | 384.5 | 364.6667 | 395.8333 | 663.1667 | 385.1667 | 388.1667 | 395.6667 | 364.5 | 402.8953 | 367.1667 | 391 | 376.8333 | 391.6667 | 399.75 | 411.3333 | 400 | 419 | 410
| Median |  | 422.44 | 423 | 404 | 365 | 376.5 | 386 | 364.5 | 395.5 | 663 | 385 | 388 | 395 | 362.5 | 402.3961 | 366.5 | 391.5 | 368.5 | 392 | 399.65 | 408 | 400 | 424 | 410
| Std Dev. |  | 8.4015 | 4.7924 | 13.8996 | 5.3448 | 2.6583 | 6.6858 | 4.4121 | 6.3061 | 3.4881 | 2.8577 | 3.8166 | 4.9666 | 12.276 | 3.7243 | 6.4936 | 6.3561 | 42.2015 | 2.582 | 2.3098 | 14.5006 | 0 | 31.029 | 10.9545
| Rel.Std.Dev. |  | 1.985% | 1.131% | 3.441% | 1.461% | 0.7045% | 1.739% | 1.21% | 1.593% | 0.526% | 0.7419% | 0.9832% | 1.255% | 3.368% | 0.9244% | 1.769% | 1.626% | 11.2% | 0.6592% | 0.5778% | 3.525% | 0% | 7.405% | 2.672%
| PDM3 |  | 7.686% | 7.861% | 2.813% | -6.9% | -3.973% | -2.149% | -7.197% | 0.735% | 68.77% | -1.98% | -1.216% | 0.6926% | -7.239% | 2.532% | -6.56% | -0.495% | -4.1% | -0.3253% | 1.732% | 4.68% | 1.795% | 6.631% | 4.34%
| Z-Score (Absolute) |  | 1.29 | 1.33 | 0.39 | 1.4 | 0.86 | 0.53 | 1.46 | 0.01 | 12.58 | 0.49 | 0.35 | 0 | 1.47 | 0.34 | 1.34 | 0.22 | 0.89 | 0.19 | 0.19 | 0.74 | 0.2 | 1.1 | 0.67
| Analytical results for Mo in OREAS 501b (Certified Value 97 ppm) | [24 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 0
1
| 1 | 1 | 108.5 | 99 | 99.6 | 90.6 | 92.1 | 99.1 | 97.9 | 101 | 5 | 84 | 97 | 96 | 89 | 92.508 | 89.6 | 101.5 | 102 | 92.5 | 98 | 119.32 | 76 | 86 |  | 1
| 1 | 2 | 109.68 | 100.8 | 98.6 | 93.2 | 87.4 | 101.8 | 100 | 103 | 5 | 88 | 99 | 96 | 88 | 94.695 | 90.6 | 101.5 | 104 | 98 | 101 | 116.51 | 75 | 86 |  | #N/A
| 1 | 3 | 107.49 | 101.6 | 104 | 91.1 | 92.8 | 99.8 | 97.9 | 103.7 | 5 | 88 | 99 | 97 | 89 | 98.524 | 92.9 | 99.9 | 111 | 94.8 | 97 | 114.39 | 65 | 87 |  | 16
| 1 | 4 | 107.09 | 101.3 | 105 | 91.9 | 90 | 100.2 | 100 | 103.9 | 5 | 87 | 100 | 97 | 85 | 97.771 | 91.3 | 102 | 136 | 95.4 | 95 | 115.53 | 75 | 87 |  | 96.8982
| 1 | 5 | 106.13 | 101.7 | 101 | 91.6 | 89.8 | 100.4 | 100 | 106.6 | 5 | 88 | 99 | 97 | 85 | 98.954 | 92.2 | 101 | 112 | 95.5 | 94 | 115.37 | 74 | 88 |  | 89
| 1 | 6 | 111.17 | 99.7 | 103 | 92.7 | 88.9 | 99.7 | 99 | 100.5 | 5 | 86 | 101 | 97 | 86 | 94.501 | 92.2 | 97.9 | 105 | 98.6 | 99 | 118.83 | 79 | 90
| Mean |  | 108.3433 | 100.6833 | 101.8667 | 91.85 | 90.1667 | 100.1667 | 99.1333 | 103.1167 | 5 | 86.8333 | 99.1667 | 96.6667 | 87 | 96.1588 | 91.4667 | 100.6333 | 111.6667 | 95.8 | 97.3333 | 116.6583 | 74 | 87.3333
| Median |  | 107.995 | 101.05 | 102 | 91.75 | 89.9 | 100 | 99.5 | 103.35 | 5 | 87.5 | 99 | 97 | 87 | 96.233 | 91.75 | 101.25 | 108 | 95.45 | 97.5 | 116.02 | 75 | 87
| Std Dev. |  | 1.8439 | 1.1017 | 2.5382 | 0.97314 | 2.0047 | 0.91797 | 1.0309 | 2.2104 | 0 | 1.6021 | 1.3292 | 0.5164 | 1.8974 | 2.6163 | 1.216 | 1.5175 | 12.5645 | 2.228 | 2.582 | 1.995 | 4.7329 | 1.5055
| Rel.Std.Dev. |  | 1.702% | 1.094% | 2.492% | 1.059% | 2.223% | 0.9164% | 1.04% | 2.144% | 0% | 1.845% | 1.34% | 0.5342% | 2.181% | 2.721% | 1.329% | 1.508% | 11.25% | 2.326% | 2.653% | 1.71% | 6.396% | 1.724%
| PDM3 |  | 11.81% | 3.906% | 5.128% | -5.21% | -6.947% | 3.373% | 2.307% | 6.418% | -94.84% | -10.39% | 2.341% | -0.2389% | -10.22% | -0.763% | -5.605% | 3.855% | 15.24% | -1.133% | 0.4491% | 20.39% | -23.63% | -9.871%
| Z-Score (Absolute) |  | 1.43 | 0.46 | 0.61 | 0.65 | 0.86 | 0.4 | 0.27 | 0.77 | 11.61 | 1.28 | 0.27 | 0.04 | 1.26 | 0.11 | 0.7 | 0.46 | 1.85 | 0.15 | 0.04 | 2.48 | 2.9 | 1.22
| Analytical results for Na in OREAS 501b (Certified Value 0.167 wt.%) | [24 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 3
3
| 1 | 1 | 0.23979 | 0.22 | 0.16 | 0.14 | 0.17 | 0.141 | 0.15 | 0.15 | 0.08 | 0.14 | 0.14 | 0.17 | 0.15 | 0.18971 | 0.14 | 0.17 | 0.11 | 0.17 | 0.2 | 0.16 | 0.21 | 0.39 | 0.21 | 1
| 1 | 2 | 0.21255 | 0.21 | 0.16 | 0.14 | 0.17 | 0.145 | 0.15 | 0.15 | 0.07 | 0.14 | 0.14 | 0.17 | 0.14 | 0.19164 | 0.15 | 0.16 | 0.11 | 0.17 | 0.21 | 0.16 | 0.2 | 0.44 | 0.21 | #N/A
| 1 | 3 | 0.2144 | 0.21 | 0.16 | 0.14 | 0.17 | 0.141 | 0.15 | 0.14 | 0.08 | 0.15 | 0.15 | 0.17 | 0.14 | 0.18621 | 0.15 | 0.16 | 0.12 | 0.16 | 0.21 | 0.16 | 0.2 | 0.38 | 0.21 | 16
| 1 | 4 | 0.2291 | 0.23 | 0.16 | 0.14 | 0.16 | 0.14 | 0.15 | 0.14 | 0.07 | 0.15 | 0.15 | 0.17 | 0.14 | 0.19022 | 0.15 | 0.17 | 0.15 | 0.16 | 0.2 | 0.17 | 0.2 | 0.41 | 0.21 | 0.16711
| 1 | 5 | 0.20698 | 0.22 | 0.14 | 0.14 | 0.17 | 0.141 | 0.15 | 0.15 | 0.08 | 0.15 | 0.15 | 0.17 | 0.14 | 0.19139 | 0.15 | 0.16 | 0.12 | 0.16 | 0.2 | 0.16 | 0.21 | 0.3 | 0.21 | 90
| 1 | 6 | 0.23662 | 0.22 | 0.15 | 0.14 | 0.16 | 0.144 | 0.15 | 0.15 | 0.07 | 0.15 | 0.15 | 0.17 | 0.14 | 0.18527 | 0.16 | 0.16 | 0.11 | 0.16 | 0.21 | 0.17 | 0.2 | 0.39 | 0.22
| Mean |  | 0.22324 | 0.21833 | 0.155 | 0.14 | 0.16667 | 0.142 | 0.15 | 0.14667 | 0.075 | 0.14667 | 0.14667 | 0.17 | 0.14167 | 0.18907 | 0.15 | 0.16333 | 0.12 | 0.16333 | 0.205 | 0.16333 | 0.20333 | 0.385 | 0.21167
| Median |  | 0.22175 | 0.22 | 0.16 | 0.14 | 0.17 | 0.141 | 0.15 | 0.15 | 0.075 | 0.15 | 0.15 | 0.17 | 0.14 | 0.18997 | 0.15 | 0.16 | 0.115 | 0.16 | 0.205 | 0.16 | 0.2 | 0.39 | 0.21
| Std Dev. |  | 0.013741 | 0.0075277 | 0.0083666 | 0 | 0.005164 | 0.002 | 0 | 0.005164 | 0.0054772 | 0.005164 | 0.005164 | 0 | 0.0040825 | 0.0026971 | 0.0063246 | 0.005164 | 0.015492 | 0.005164 | 0.0054772 | 0.005164 | 0.005164 | 0.046797 | 0.0040825
| Rel.Std.Dev. |  | 6.155% | 3.448% | 5.398% | 0% | 3.098% | 1.408% | 0% | 3.521% | 7.303% | 3.521% | 3.521% | 0% | 2.882% | 1.426% | 4.216% | 3.162% | 12.91% | 3.162% | 2.672% | 3.162% | 2.54% | 12.16% | 1.929%
| PDM3 |  | 33.59% | 30.65% | -7.247% | -16.22% | -0.2654% | -15.03% | -10.24% | -12.23% | -55.12% | -12.23% | -12.23% | 1.729% | -15.23% | 13.14% | -10.24% | -2.26% | -28.19% | -2.26% | 22.67% | -2.26% | 21.68% | 130.4% | 26.66%
| Z-Score (Absolute) |  | 1.89 | 1.74 | 0.26 | 0.74 | 0.11 | 0.67 | 0.42 | 0.53 | 2.79 | 0.53 | 0.53 | 0.21 | 0.68 | 0.81 | 0.42 | 0 | 1.37 | 0 | 1.32 | 0 | 1.26 | 7.01 | 1.53
| Analytical results for Nb in OREAS 501b (Indicative Value 1.35 ppm) | [24 empty] | Indicative
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 | [6 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | [6 empty] | 2
2
| 1 | 1 | 2.12 | 0.8 | 1.22 | 1.16 | 1.65 | 1.5 | < 1 | 1.35 | 2.7304 | 0.96 | 1.29 | 8 | 1.01 | 1 | 1.64 | 30 | < 5 | [6 empty] | 1
| 1 | 2 | 2.25 | 0.8 | 1.27 | 1.05 | 1.67 | 1.37 | < 1 | 1.56 | 3.0187 | 0.86 | 1.18 | 8 | 1.12 | 1 | 1.94 | 29 | < 5 | [6 empty] | 10
| 1 | 3 | 2.29 | 0.8 | 1.11 | 1.16 | 1.75 | 1.43 | < 1 | 1.64 | 3.1979 | 0.78 | 1.28 | 8 | 1.11 | 1 | 1.98 | 25 | < 5 | [6 empty] | 16
| 1 | 4 | 2.34 | 0.8 | 1.18 | 1.06 | 1.9 | 1.43 | < 1 | 1.41 | 3.1172 | 0.94 | 1.39 | 10 | 1.17 | 1 | 1.82 | 31 | < 5 | [6 empty] | 1.3469
| 1 | 5 | 2.43 | 0.6 | 1.25 | 1.03 | 1.85 | 1.67 | < 1 | 1.58 | 2.86 | 0.9 | 1.35 | 9 | 1.13 | 1 | 1.83 | 30 | < 5 | [6 empty] | 16
| 1 | 6 | 2.47 | 0.8 | 1.12 | 1.05 | 1.77 | 1.26 | < 1 | 1.45 | 2.5053 | 0.73 | 1.07 | 9 | 1.02 | 1 | 1.88 | 26 | < 5
| Mean |  | 2.3167 | 0.76667 | 1.1917 | 1.085 | 1.765 | 1.4433 |  | 1.4983 | 2.9049 | 0.86167 | 1.26 | 8.6667 | 1.0933 | 1 | 1.8483 | 28.5
| Median |  | 2.315 | 0.8 | 1.2 | 1.055 | 1.76 | 1.43 |  | 1.505 | 2.9393 | 0.88 | 1.285 | 8.5 | 1.115 | 1 | 1.855 | 29.5
| Std Dev. |  | 0.12707 | 0.08165 | 0.066758 | 0.058907 | 0.097929 | 0.13706 |  | 0.11197 | 0.25922 | 0.090866 | 0.1173 | 0.8165 | 0.064083 | 0 | 0.1194 | 2.429
| Rel.Std.Dev. |  | 5.485% | 10.65% | 5.602% | 5.429% | 5.548% | 9.496% |  | 7.473% | 8.923% | 10.55% | 9.31% | 9.421% | 5.861% | 0% | 6.46% | 8.523%
| PDM3 |  | 71.99% | -43.08% | -11.53% | -19.45% | 31.04% | 7.156% |  | 11.24% | 115.7% | -36.03% | -6.455% | 543.4% | -18.83% | -25.76% | 37.22% | 2016%
| Z-Score (Absolute) |  | 1.33 | 1.03 | 0.38 | 0.55 | 0.49 | 0 | 0.67 | 0.08 | 2.22 | 0.88 | 0.28 | 10.99 | 0.53 | 0.67 | 0.62 | 41.15 | 5.41
| Analytical results for Nd in OREAS 501b (Certified Value 26.2 ppm) | [24 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS | 1
2
| 1 | 1 | 01: 28.323; 02: 26.63; 03: 25.7; 04: 23.65; 05: 25.6 | 1
| 1 | 2 | 01: 28.896; 02: 24.95; 03: 25.2; 04: 25.71; 05: 25.6 | #N/A
| 1 | 3 | 01: 28.511; 02: 26.33; 03: 25.7; 04: 24.19; 05: 27.1 | 16
| 1 | 4 | 01: 27.933; 02: 25.57; 03: 26.2; 04: 25.22; 05: 27.5 | 26.181
| 1 | 5 | 01: 27.807; 02: 25.25; 03: 24.3; 04: 25.47; 05: 27.2 | 91
| 1 | 6 | 01: 28.699; 02: 25.34; 03: 25.5; 04: 24.09; 05: 25.9
| Mean |  | 01: 28.3615; 02: 25.6783; 03: 25.4333; 04: 24.7217; 05: 26.4833
| Median |  | 01: 28.417; 02: 25.455; 03: 25.6; 04: 24.705; 05: 26.5
| Std Dev. |  | 01: 0.42767; 02: 0.65886; 03: 0.64395; 04: 0.85033; 05: 0.87502
| Rel.Std.Dev. |  | 01: 1.508%; 02: 2.566%; 03: 2.532%; 04: 3.44%; 05: 3.304%
| PDM3 |  | 01: 8.329%; 02: -1.92%; 03: -2.856%; 04: -5.574%; 05: 1.155%
| Z-Score (Absolute) |  | 01: 2.25; 02: 0; 03: 0.21; 04: 0.8; 05: 0.67
| Analytical results for Ni in OREAS 501b (Certified Value 37.5 ppm) | [24 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 1
2
| 1 | 1 | 41.792 | 39.4 | 41 | 38.1 | 38.9 | 36.7 | 37.6 | 37.4 | 10 | 37 | 39 | 37.5 | 36 | 36.6772 | 35.4 | 38.1 | 32 | 38.4 | 38 | 45.9 | 34 | 36 |  | 1
| 1 | 2 | 45.263 | 38.1 | 38 | 38.1 | 41.9 | 39.2 | 40.1 | 38.5 | 10 | 36 | 39 | 36.3 | 38 | 36.2438 | 35.2 | 40.8 | 32 | 41.2 | 37 | 47 | 34 | 34 |  | #N/A
| 1 | 3 | 46.192 | 39.2 | 41 | 34.7 | 45.5 | 36.5 | 40.2 | 37.9 | 10 | 36 | 39 | 37.8 | 34 | 36.4417 | 35.1 | 38.6 | 34 | 38.7 | 36 | 40.6 | 33 | 34 |  | 16
| 1 | 4 | 43.147 | 40.6 | 43 | 36.9 | 40.4 | 38.8 | 39.2 | 36.7 | 10 | 38 | 39 | 36 | 33 | 37.2256 | 35.2 | 39.4 | 40 | 40.3 | 37 | 45.9 | 37 | 37 |  | 37.4761
| 1 | 5 | 44.784 | 40.5 | 39 | 37.7 | 39.1 | 38.4 | 38.1 | 41.5 | 10 | 37 | 39 | 38.8 | 31 | 39.0197 | 37.6 | 37.8 | 35 | 39 | 36 | 45.2 | 37 | 35 |  | 92
| 1 | 6 | 45.187 | 39.2 | 38 | 38.3 | 38.4 | 37.7 | 38.3 | 38.2 | 10 | 36 | 40 | 38.2 | 31 | 36.2297 | 38.6 | 36 | 34 | 40.2 | 38 | 45.6 | 36 | 36
| Mean |  | 44.3942 | 39.5 | 40 | 37.3 | 40.7 | 37.8833 | 38.9167 | 38.3667 | 10 | 36.6667 | 39.1667 | 37.4333 | 33.8333 | 36.9729 | 36.1833 | 38.45 | 34.5 | 39.6333 | 37 | 45.0333 | 35.1667 | 35.3333
| Median |  | 44.9855 | 39.3 | 40 | 37.9 | 39.75 | 38.05 | 38.75 | 38.05 | 10 | 36.5 | 39 | 37.65 | 33.5 | 36.5594 | 35.3 | 38.35 | 34 | 39.6 | 37 | 45.75 | 35 | 35.5
| Std Dev. |  | 1.6183 | 0.93381 | 2 | 1.3682 | 2.6706 | 1.1125 | 1.087 | 1.6609 | 0 | 0.8165 | 0.40825 | 1.0893 | 2.7869 | 1.0684 | 1.5211 | 1.6121 | 2.9496 | 1.0967 | 0.89443 | 2.2527 | 1.7224 | 1.2111
| Rel.Std.Dev. |  | 3.645% | 2.364% | 5% | 3.668% | 6.562% | 2.937% | 2.793% | 4.329% | 0% | 2.227% | 1.042% | 2.91% | 8.237% | 2.89% | 4.204% | 4.193% | 8.549% | 2.767% | 2.417% | 5.002% | 4.898% | 3.428%
| PDM3 |  | 18.46% | 5.4% | 6.735% | -0.47% | 8.602% | 1.087% | 3.844% | 2.376% | -73.32% | -2.16% | 4.511% | -0.1142% | -9.72% | -1.343% | -3.45% | 2.599% | -7.941% | 5.756% | -1.27% | 20.17% | -6.162% | -5.718%
| Z-Score (Absolute) |  | 2.71 | 0.74 | 0.94 | 0.14 | 1.22 | 0.09 | 0.51 | 0.29 | 11.13 | 0.4 | 0.61 | 0.09 | 1.54 | 0.28 | 0.59 | 0.32 | 1.27 | 0.8 | 0.27 | 2.97 | 1 | 0.94
| Analytical results for P in OREAS 501b (Certified Value 0.1 wt.%) | [24 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 24 | 25 | 26 | 27 |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  | 3
3
| 1 | 1 | 0.098774 | 0.095 | 0.098 | 0.099 | 0.102 | 0.107 | 0.104 | 0.0932 | 0.06 | 0.099 | 0.102 | 0.092 | 0.098309 | 0.096 | 0.102 | 0.0872 | 0.101 | 0.099 | 0.0995 | 0.14 | 0.094 |  |  | 1
| 1 | 2 | 0.0969 | 0.089 | 0.098 | 0.099 | 0.1 | 0.11 | 0.106 | 0.0973 | 0.06 | 0.1 | 0.103 | 0.092 | 0.10553 | 0.097 | 0.1 | 0.086 | 0.111 | 0.105 | 0.1037 | 0.09 | 0.096 |  |  | #N/A
| 1 | 3 | 0.10193 | 0.094 | 0.1 | 0.096 | 0.102 | 0.107 | 0.107 | 0.0943 | 0.06 | 0.101 | 0.104 | 0.092 | 0.1015 | 0.1 | 0.099 | 0.0935 | 0.104 | 0.104 | 0.1056 | 0.1 | 0.094 |  |  | 16
| 1 | 4 | 0.099718 | 0.096 | 0.102 | 0.098 | 0.102 | 0.1 | 0.108 | 0.0926 | 0.06 | 0.103 | 0.105 | 0.089 | 0.10376 | 0.098 | 0.103 | 0.124 | 0.103 | 0.109 | 0.1035 | 0.1 | 0.094 |  |  | 0.099658
| 1 | 5 | 0.10399 | 0.093 | 0.098 | 0.098 | 0.101 | 0.104 | 0.108 | 0.0975 | 0.06 | 0.1 | 0.104 | 0.089 | 0.10923 | 0.099 | 0.1 | 0.1006 | 0.104 | 0.106 | 0.101 | 0.1 | 0.096 |  |  | 93
| 1 | 6 | 0.10596 | 0.089 | 0.094 | 0.1 | 0.101 | 0.106 | 0.107 | 0.0941 | 0.06 | 0.103 | 0.103 | 0.09 | 0.10293 | 0.101 | 0.099 | 0.093 | 0.105 | 0.108 | 0.1014 | 0.1 | 0.098
| Mean |  | 0.10121 | 0.092667 | 0.098333 | 0.098333 | 0.10133 | 0.10567 | 0.10667 | 0.094833 | 0.06 | 0.101 | 0.1035 | 0.090667 | 0.10354 | 0.0985 | 0.1005 | 0.097383 | 0.10467 | 0.10517 | 0.10245 | 0.105 | 0.095333
| Median |  | 0.10082 | 0.0935 | 0.098 | 0.0985 | 0.1015 | 0.1065 | 0.107 | 0.0942 | 0.06 | 0.1005 | 0.1035 | 0.091 | 0.10335 | 0.0985 | 0.1 | 0.09325 | 0.104 | 0.1055 | 0.10245 | 0.1 | 0.095
| Std Dev. |  | 0.0033925 | 0.0030111 | 0.0026583 | 0.0013663 | 0.0008165 | 0.0033862 | 0.0015055 | 0.002082 | 0 | 0.0016733 | 0.0010488 | 0.0015055 | 0.0036956 | 0.0018708 | 0.0016432 | 0.014045 | 0.0033862 | 0.0035449 | 0.0022134 | 0.017607 | 0.001633
| Rel.Std.Dev. |  | 3.352% | 3.249% | 2.703% | 1.389% | 0.8058% | 3.205% | 1.411% | 2.195% | 0% | 1.657% | 1.013% | 1.661% | 3.569% | 1.899% | 1.635% | 14.42% | 3.235% | 3.371% | 2.16% | 16.77% | 1.713%
| PDM3 |  | 1.559% | -7.016% | -1.33% | -1.33% | 1.681% | 6.029% | 7.032% | -4.842% | -39.79% | 1.346% | 3.855% | -9.022% | 3.9% | -1.162% | 0.8446% | -2.283% | 5.026% | 5.527% | 2.801% | 5.36% | -4.34%
| Z-Score (Absolute) |  | 0.04 | 1.55 | 0.5 | 0.5 | 0.06 | 0.87 | 1.06 | 1.15 | 7.64 | 0 | 0.47 | 1.93 | 0.47 | 0.47 | 0.09 | 0.67 | 0.68 | 0.78 | 0.27 | 0.75 | 1.06
| Analytical results for Pb in OREAS 501b (Certified Value 9.42 ppm) | [24 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | AR*OES |  | 2
3
| 1 | 1 | 10.233 | 9.5 | 10 | 9.2 | 9 | 9.3 | 9 | 10.3 | 18 | 2 | 14 | 9.8 | 10 | 8.95 | 9 | 10.3 | 14 | 11.5 | 11 | 9.9 | 14 | 29 |  | 1
| 1 | 2 | 10.679 | 9.3 | 10 | 9.4 | 9.2 | 9.3 | 9.5 | 10.1 | 17 | 3 | 10 | 9.2 | 9 | 8.97 | 8.5 | 10 | 11 | 9.7 | 10 | 10 | 14 | 28 |  | #N/A
| 1 | 3 | 10.156 | 9.8 | 10 | 8.8 | 9.3 | 8.8 | 9.1 | 9.9 | 18 | 3 | 7 | 8.7 | 9 | 9.55 | 9.1 | 9.7 | 14 | 9.3 | 9 | 11.1 | 14 | 29 |  | 16
| 1 | 4 | 10.029 | 9.8 | 10 | 9.1 | 9.8 | 8.9 | 9.4 | 10.6 | 17 | 2 | 8 | 9.2 | 9 | 9.75 | 8.8 | 9.8 | 17 | 9 | 9 | 9.4 | 15 | 29 |  | 9.418
| 1 | 5 | 9.9632 | 9.7 | 10 | 9 | 9.5 | 9.4 | 9.1 | 12.2 | 19 | 3 | 8 | 9.1 | 8 | 9.26 | 9 | 9.6 | 13 | 8.6 | 9 | 10.5 | 14 | 29 |  | 94
| 1 | 6 | 10.702 | 9.3 | 9 | 8.9 | 9.7 | 9 | 9.2 | 9.8 | 19 | 2 | 8 | 9.2 | 8 | 9.01 | 9 | 9.5 | 11 | 8.8 | 10 | 9.6 | 15 | 30
| Mean |  | 10.2937 | 9.5667 | 9.8333 | 9.0667 | 9.4167 | 9.1167 | 9.2167 | 10.4833 | 18 | 2.5 | 9.1667 | 9.2 | 8.8333 | 9.2483 | 8.9 | 9.8167 | 13.3333 | 9.4833 | 9.6667 | 10.0833 | 14.3333 | 29
| Median |  | 10.1945 | 9.6 | 10 | 9.05 | 9.4 | 9.15 | 9.15 | 10.2 | 18 | 2.5 | 8 | 9.2 | 9 | 9.135 | 9 | 9.75 | 13.5 | 9.15 | 9.5 | 9.95 | 14 | 29
| Std Dev. |  | 0.32159 | 0.23381 | 0.40825 | 0.21602 | 0.30605 | 0.24833 | 0.19408 | 0.88863 | 0.89443 | 0.54772 | 2.5626 | 0.35214 | 0.75277 | 0.33648 | 0.21909 | 0.29269 | 2.2509 | 1.061 | 0.8165 | 0.62423 | 0.5164 | 0.63246
| Rel.Std.Dev. |  | 3.124% | 2.444% | 4.152% | 2.383% | 3.25% | 2.724% | 2.106% | 8.477% | 4.969% | 21.91% | 27.96% | 3.828% | 8.522% | 3.638% | 2.462% | 2.982% | 16.88% | 11.19% | 8.447% | 6.191% | 3.603% | 2.181%
| PDM3 |  | 9.299% | 1.579% | 4.41% | -3.73% | -0.01376% | -3.199% | -2.137% | 11.31% | 91.12% | -73.45% | -2.668% | -2.314% | -6.208% | -1.801% | -5.5% | 4.233% | 41.57% | 0.6941% | 2.641% | 7.065% | 52.19% | 207.9%
| Z-Score (Absolute) |  | 1.2 | 0.06 | 0.48 | 0.71 | 0.17 | 0.64 | 0.48 | 1.49 | 13.19 | 10.93 | 0.56 | 0.51 | 1.08 | 0.43 | 0.97 | 0.45 | 5.93 | 0.06 | 0.22 | 0.87 | 7.48 | 30.3
| Analytical results for Pd in OREAS 501b (Indicative Value < 10 ppb) | [24 empty] | Indicative
Pd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab |  | 1
| No | No | 01: 02; 02: 03; 03: 08; 04: 25 |  | ppb
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*OES/MS |  | 1
1
| 1 | 1 | 01: 10; 02: < 10; 03: < 10; 04: < 100 |  | 1
| 1 | 2 | 01: 10; 02: < 10; 03: < 10; 04: < 100 |  | 1
| 1 | 3 | 01: 10; 02: < 10; 03: < 10; 04: < 100 |  | 16
| 1 | 4 | 01: < 10; 02: < 10; 03: < 10; 04: < 100 |  | < 10
| 1 | 5 | 01: 10; 02: < 10; 03: < 10; 04: < 100 |  | 17
| 1 | 6 | 01: 10; 02: < 10; 03: < 10; 04: < 100
| Mean |  | 01: 10
| Median |  | 01: 10
| Std Dev. |  | 01: 0
Rel.Std.Dev. | bar: 01
PDM3
Z-Score (Absolute) | bar: 01 to 04
| Analytical results for Pr in OREAS 501b (Certified Value 7.17 ppm) | [24 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS |  | 2
3
| 1 | 1 | 01: 7.66; 02: 6.883; 03: 6.92; 04: 6.931; 05: 6.9 |  | 1
| 1 | 2 | 01: 7.8359; 02: 4.884; 03: 6.99; 04: 7.599; 05: 6.9 |  | #N/A
| 1 | 3 | 01: 7.6328; 02: 7.063; 03: 7.06; 04: 7.183; 05: 7.2 |  | 16
| 1 | 4 | 01: 7.4888; 02: 6.866; 03: 7.02; 04: 7.289; 05: 7.3 |  | 7.173
| 1 | 5 | 01: 7.4313; 02: 6.99; 03: 6.71; 04: 7.416; 05: 7.5 |  | 95
| 1 | 6 | 01: 7.7628; 02: 6.909; 03: 6.84; 04: 6.966; 05: 7
| Mean |  | 01: 7.6353; 02: 6.5992; 03: 6.9233; 04: 7.2307; 05: 7.1333
| Median |  | 01: 7.6464; 02: 6.896; 03: 6.955; 04: 7.236; 05: 7.1
| Std Dev. |  | 01: 0.15504; 02: 0.8435; 03: 0.13033; 04: 0.25905; 05: 0.24221
| Rel.Std.Dev. |  | 01: 2.031%; 02: 12.78%; 03: 1.883%; 04: 3.583%; 05: 3.395%
| PDM3 |  | 01: 6.445%; 02: -7.999%; 03: -3.48%; 04: 0.8045%; 05: -0.5524%
| Z-Score (Absolute) |  | 01: 1.61; 02: 1.72; 03: 0.67; 04: 0.31; 05: 0
| Analytical results for Pt in OREAS 501b (Indicative Value < 5 ppb) | [24 empty] | Indicative
Pt | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 08; 04: 25 |  | ppb
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*OES/MS |  | 2
2
| 1 | 1 | 01: 6.9711; 02: < 5; 03: < 5; 04: < 100 |  | 1
| 1 | 2 | 01: 7.0988; 02: < 5; 03: < 5; 04: < 100 |  | 1
| 1 | 3 | 01: 5.5393; 02: < 5; 03: < 5; 04: < 100 |  | 16
| 1 | 4 | 01: 5.716; 02: < 5; 03: < 5; 04: < 100 |  | < 5
| 1 | 5 | 01: 5.7013; 02: < 5; 03: < 5; 04: < 100 |  | 13
| 1 | 6 | 01: 5.8435; 02: < 5; 03: < 5; 04: < 100
| Mean |  | 01: 6.145
| Median |  | 01: 5.7797
| Std Dev. |  | 01: 0.69726
| Rel.Std.Dev. |  | 01: 11.35%
PDM3
| Z-Score (Absolute) |  | 01: 0.67; 02: 0.67; 03: 0.67; 04: 111.22
| Analytical results for Rb in OREAS 501b (Certified Value 126 ppm) | [24 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | [8 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [8 empty] | 0
0
| 1 | 1 | 129.44 | 126.09 | 147 | 121 | 126 | 116 | 127.64 | 124 | 115.3261 | 120.5 | 130 | 124 | 123 | 130.5 | 147.1 | [8 empty] | 1
| 1 | 2 | 133.9 | 126.32 | 145 | 124.5 | 132.5 | 122 | 130.13 | 119 | 116.5373 | 120 | 133 | 115 | 130 | 133.5 | 144.3 | [8 empty] | #N/A
| 1 | 3 | 131.88 | 127.25 | 153 | 122 | 131 | 121 | 128.15 | 124 | 120.0698 | 122.5 | 125 | 133 | 129.5 | 134 | 144.5 | [8 empty] | 16
| 1 | 4 | 127.4 | 126.85 | 156 | 118 | 128.5 | 124 | 128.12 | 123 | 119.9736 | 123.5 | 129.5 | 151 | 127 | 134.3 | 139.7 | [8 empty] | 125.8779
| 1 | 5 | 127.74 | 127.37 | 147 | 120 | 132.5 | 121 | 132.17 | 120.5 | 119.1317 | 132.5 | 127 | 132 | 126 | 134.5 | 144.3 | [8 empty] | 96
| 1 | 6 | 133.25 | 127.96 | 144 | 123.5 | 129 | 125 | 128.88 | 124.5 | 116.6955 | 124 | 121.5 | 127 | 127.5 | 132.1 | 145.2
| Mean |  | 130.6017 | 126.9733 | 148.6667 | 121.5 | 129.9167 | 121.5 | 129.1817 | 122.5 | 117.9557 | 123.8333 | 127.6667 | 130.3333 | 127.1667 | 133.15 | 144.1833
| Median |  | 130.66 | 127.05 | 147 | 121.5 | 130 | 121.5 | 128.515 | 123.5 | 117.9136 | 123 | 128.25 | 129.5 | 127.25 | 133.75 | 144.4
| Std Dev. |  | 2.8041 | 0.6971 | 4.761 | 2.3664 | 2.5577 | 3.1464 | 1.7023 | 2.2361 | 2.0218 | 4.535 | 4.0702 | 12.0277 | 2.543 | 1.5566 | 2.44
| Rel.Std.Dev. |  | 2.147% | 0.549% | 3.202% | 1.948% | 1.969% | 2.59% | 1.318% | 1.825% | 1.714% | 3.662% | 3.188% | 9.228% | 2% | 1.169% | 1.692%
| PDM3 |  | 3.753% | 0.8703% | 18.1% | -3.478% | 3.209% | -3.478% | 2.625% | -2.683% | -6.294% | -1.624% | 1.421% | 3.54% | 1.024% | 5.777% | 14.54%
| Z-Score (Absolute) |  | 0.52 | 0.12 | 3.69 | 1.08 | 0.4 | 1.08 | 0.27 | 0.91 | 1.71 | 0.67 | 0 | 0.47 | 0.09 | 0.96 | 2.91
| Analytical results for Re in OREAS 501b (Certified Value 0.003 ppm) | [24 empty] | Certified
Re | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab |  | 1
| No | No | 01: 02; 02: 04; 03: 05; 04: 08; 05: 12; 06: 17; 07: 18; 08: 20; 09: 23; 10: 25 |  | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*MS; 04: AR*MS; 05: AR*MS; 06: AR*MS; 07: AR*MS; 08: AR*MS; 09: AR*OES/MS; 10: AR*OES/MS |  | 3
3
| 1 | 1 | 01: 0.003; 02: 0.003; 03: 0.003; 04: 0.002; 05: 0.003; 06: 0.003; 07: 0.003; 08: 0.003; 09: < 0.01; 10: < 0.1 |  | 1
| 1 | 2 | 01: 0.003; 02: 0.003; 03: 0.003; 04: 0.004; 05: 0.002; 06: 0.003; 07: 0.003; 08: 0.003; 09: < 0.01; 10: < 0.1 |  | 4
| 1 | 3 | 01: 0.003; 02: 0.003; 03: 0.003; 04: 0.004; 05: 0.003; 06: 0.003; 07: 0.003; 08: 0.003; 09: < 0.01; 10: < 0.1 |  | 16
| 1 | 4 | 01: 0.004; 02: 0.003; 03: 0.003; 04: 0.005; 05: 0.002; 06: 0.003; 07: 0.003; 08: 0.002; 09: 0.01; 10: < 0.1 |  | 0.0029048
| 1 | 5 | 01: 0.004; 02: 0.003; 03: 0.003; 04: 0.004; 05: 0.003; 06: 0.003; 07: 0.003; 08: 0.003; 09: < 0.01; 10: < 0.1 |  | 97
| 1 | 6 | 01: 0.002; 02: 0.002; 03: 0.002; 04: 0.004; 05: 0.002; 06: 0.003; 07: 0.003; 08: 0.002; 09: < 0.01; 10: < 0.1
| Mean |  | 01: 0.0031667; 02: 0.0028333; 03: 0.0028333; 04: 0.0038333; 05: 0.0025; 06: 0.003; 07: 0.003; 08: 0.0026667; 09: 0.01
| Median |  | 01: 0.003; 02: 0.003; 03: 0.003; 04: 0.004; 05: 0.0025; 06: 0.003; 07: 0.003; 08: 0.003; 09: 0.01
| Std Dev. |  | 01: 0.00075277; 02: 0.00040825; 03: 0.00040825; 04: 0.00098319; 05: 0.00054772; 06: 4.7507e-19; 07: 4.7507e-19; 08: 0.0005164
| Rel.Std.Dev. |  | 01: 23.77%; 02: 14.41%; 03: 14.41%; 04: 25.65%; 05: 21.91%; 06: 1.584e-14%; 07: 1.584e-14%; 08: 19.36%
| PDM3 |  | 01: 9.016%; 02: -2.459%; 03: -2.459%; 04: 31.97%; 05: -13.93%; 06: 3.279%; 07: 3.279%; 08: -8.197%; 09: 244.3%
| Z-Score (Absolute) |  | 01: 0.45; 02: 0.45; 03: 0.45; 04: 2.25; 05: 1.35; 06: 0; 07: 0; 08: 0.9; 09: 18.88; 10: 261.63
| Analytical results for S in OREAS 501b (Certified Value 0.364 wt.%) | [24 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 24 | 25 | 26 | 27 |  |  | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  | 3
3
| 1 | 1 | 0.41378 | 0.349 | 0.35 | 0.36 | 0.39 | 0.32 | 0.31 | 0.377 | 2.71 | 0.32 | 0.38 | 0.35 | 0.373 | 0.35 | 0.37 | 0.41 | 0.36 | 0.4 | 0.35 | 0.39 | 0.32 |  |  | 1
| 1 | 2 | 0.38472 | 0.348 | 0.355 | 0.36 | 0.38 | 0.35 | 0.33 | 0.3897 | 2.41 | 0.34 | 0.38 | 0.34 | 0.392 | 0.35 | 0.36 | 0.4 | 0.4 | 0.4 | 0.36 | 0.36 | 0.33 |  |  | #N/A
| 1 | 3 | 0.39628 | 0.356 | 0.36 | 0.35 | 0.38 | 0.34 | 0.32 | 0.3816 | 2.05 | 0.34 | 0.38 | 0.35 | 0.386 | 0.36 | 0.36 | 0.43 | 0.38 | 0.4 | 0.36 | 0.34 | 0.33 |  |  | 16
| 1 | 4 | 0.40651 | 0.361 | 0.36 | 0.36 | 0.38 | 0.33 | 0.32 | 0.3793 | 1.82 | 0.35 | 0.39 | 0.33 | 0.392 | 0.36 | 0.37 | 0.52 | 0.38 | 0.42 | 0.35 | 0.35 | 0.31 |  |  | 0.36421
| 1 | 5 | 0.39084 | 0.356 | 0.36 | 0.36 | 0.38 | 0.34 | 0.33 | 0.3758 | 1.72 | 0.34 | 0.39 | 0.33 | 0.415 | 0.35 | 0.36 | 0.43 | 0.37 | 0.4 | 0.35 | 0.34 | 0.31 |  |  | 98
| 1 | 6 | 0.4288 | 0.349 | 0.355 | 0.36 | 0.38 | 0.34 | 0.32 | 0.3634 | 1.45 | 0.35 | 0.38 | 0.33 | 0.374 | 0.36 | 0.36 | 0.4 | 0.37 | 0.42 | 0.35 | 0.35 | 0.32
| Mean |  | 0.40349 | 0.35317 | 0.35667 | 0.35833 | 0.38167 | 0.33667 | 0.32167 | 0.3778 | 2.0267 | 0.34 | 0.38333 | 0.33833 | 0.38867 | 0.355 | 0.36333 | 0.43167 | 0.37667 | 0.40667 | 0.35333 | 0.355 | 0.32
| Median |  | 0.4014 | 0.3525 | 0.3575 | 0.36 | 0.38 | 0.34 | 0.32 | 0.37815 | 1.935 | 0.34 | 0.38 | 0.335 | 0.389 | 0.355 | 0.36 | 0.42 | 0.375 | 0.4 | 0.35 | 0.35 | 0.32
| Std Dev. |  | 0.016248 | 0.0052694 | 0.0040825 | 0.0040825 | 0.0040825 | 0.010328 | 0.0075277 | 0.0086058 | 0.46556 | 0.010954 | 0.005164 | 0.0098319 | 0.015384 | 0.0054772 | 0.005164 | 0.04535 | 0.013663 | 0.010328 | 0.005164 | 0.018708 | 0.0089443
| Rel.Std.Dev. |  | 4.027% | 1.492% | 1.145% | 1.139% | 1.07% | 3.068% | 2.34% | 2.278% | 22.97% | 3.222% | 1.347% | 2.906% | 3.958% | 1.543% | 1.421% | 10.51% | 3.627% | 2.54% | 1.462% | 5.27% | 2.795%
| PDM3 |  | 10.79% | -3.031% | -2.07% | -1.612% | 4.794% | -7.561% | -11.68% | 3.732% | 456.5% | -6.646% | 5.252% | -7.104% | 6.716% | -2.528% | -0.2396% | 18.52% | 3.421% | 11.66% | -2.985% | -2.528% | -12.14%
| Z-Score (Absolute) |  | 1.52 | 0.17 | 0.06 | 0 | 0.79 | 0.73 | 1.24 | 0.66 | 56.25 | 0.62 | 0.84 | 0.67 | 1.02 | 0.11 | 0.17 | 2.47 | 0.62 | 1.63 | 0.17 | 0.11 | 1.29
| Analytical results for Sb in OREAS 501b (Certified Value 0.49 ppm) | [24 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 2
2
| 1 | 1 | 0.4218 | 0.39 | 0.54 | 0.48 | 0.48 | 0.5 | 0.35 | 0.41 | < 5 | < 1 | < 3 | 0.61 | 5 | 0.531 | 0.54 | 0.47 | < 5 | 0.51 | < 1 | 0.5 | 0.57 | < 2 | < 5 | 1
| 1 | 2 | 0.4794 | 0.38 | 0.52 | 0.5 | 0.52 | 0.4 | 0.34 | 0.38 | < 5 | < 1 | < 3 | 0.6 | 2 | 0.451 | 0.51 | 0.51 | < 5 | 0.5 | < 1 | 0.5 | 0.61 | < 2 | < 5 | #N/A
| 1 | 3 | 0.4262 | 0.43 | 0.54 | 0.51 | 0.49 | 0.5 | 0.33 | 0.37 | < 5 | < 1 | < 3 | 0.65 | < 2 | 0.629 | 0.52 | 0.47 | < 5 | 0.48 | < 1 | 0.5 | 0.59 | < 2 | < 5 | 16
| 1 | 4 | 0.4585 | 0.4 | 0.56 | 0.5 | 0.52 | 0.4 | 0.37 | 0.39 | < 5 | < 1 | < 3 | 0.62 | < 2 | 0.494 | 0.53 | 0.47 | < 5 | 0.5 | < 1 | 0.5 | 0.61 | < 2 | < 5 | 0.4856
| 1 | 5 | 0.4234 | 0.38 | 0.54 | 0.69 | 0.54 | 0.5 | 0.35 | 0.41 | < 5 | < 1 | < 3 | 0.67 | < 2 | 0.477 | 0.51 | 0.44 | < 5 | 0.48 | < 1 | 0.5 | 0.61 | < 2 | < 5 | 99
| 1 | 6 | 0.4585 | 0.37 | 0.52 | 0.51 | 0.48 | 0.5 | 0.37 | 0.41 | < 5 | < 1 | < 3 | 0.67 | < 2 | 0.443 | 0.51 | 0.5 | < 5 | 0.58 | < 1 | 0.5 | 0.63 | < 2 | < 5
| Mean |  | 0.44463 | 0.39167 | 0.53667 | 0.53167 | 0.505 | 0.46667 | 0.35167 | 0.395 |  |  |  | 0.63667 | 3.5 | 0.50417 | 0.52 | 0.47667 |  | 0.50833 |  | 0.5 | 0.60333
| Median |  | 0.44235 | 0.385 | 0.54 | 0.505 | 0.505 | 0.5 | 0.35 | 0.4 |  |  |  | 0.635 | 3.5 | 0.4855 | 0.515 | 0.47 |  | 0.5 |  | 0.5 | 0.61
| Std Dev. |  | 0.024105 | 0.02137 | 0.015055 | 0.078337 | 0.0251 | 0.05164 | 0.016021 | 0.017607 |  |  |  | 0.030768 | 2.1213 | 0.068866 | 0.012649 | 0.025033 |  | 0.037103 |  | 0 | 0.020656
| Rel.Std.Dev. |  | 5.421% | 5.456% | 2.805% | 14.73% | 4.97% | 11.07% | 4.556% | 4.457% |  |  |  | 4.833% | 60.61% | 13.66% | 2.433% | 5.252% |  | 7.299% |  | 0% | 3.424%
| PDM3 |  | -8.437% | -19.34% | 10.52% | 9.486% | 3.995% | -3.899% | -27.58% | -18.66% |  |  |  | 31.11% | 620.8% | 3.823% | 7.083% | -1.84% |  | 4.681% |  | 2.965% | 24.24%
| Z-Score (Absolute) |  | 0.56 | 0.9 | 0.03 | 0 | 0.17 | 0.42 | 1.16 | 0.88 | 28.7 | 3.01 | 15.85 | 0.67 | 12.64 | 0.18 | 0.07 | 0.35 | 28.7 | 0.15 | 3.01 | 0.2 | 0.46 | 9.43 | 28.7
| Analytical results for Sc in OREAS 501b (Certified Value 7.08 ppm) | [24 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  | 2
2
| 1 | 1 | 8.0315 | 7.3 | 6 | 7 | 8.3 | 6.7 | 7.9 | 7.3 | 6.8 | 6 | 5 | 7.7 | 6 | 7.8761 | 6.5 | 7.5 | 6 | 7.1 | 7 | 8.8 | 8 | 6 |  | 1
| 1 | 2 | 8.1026 | 7.4 | 6 | 7.1 | 8.2 | 6.6 | 8.5 | 7.7 | 6.8 | 6 | 5 | 7.2 | 6 | 7.9959 | 6.6 | 7.6 | 6 | 7.3 | 7 | 8.5 | 9 | 6 |  | #N/A
| 1 | 3 | 7.7872 | 7.4 | 6 | 7 | 8.6 | 6.4 | 8.2 | 7.4 | 6.8 | 6 | 5 | 7.3 | 6 | 8.1062 | 6.6 | 7 | 6 | 7.2 | 7 | 8.4 | 9 | 6 |  | 16
| 1 | 4 | 7.9279 | 7.3 | 6 | 7.1 | 8.2 | 6.1 | 8.3 | 7.4 | 6.8 | 6 | 5 | 7.3 | 6 | 8.2051 | 6.6 | 7.7 | 8 | 7.4 | 7 | 8.5 | 9 | 6 |  | 7.0818
| 1 | 5 | 7.7215 | 7.3 | 6 | 7.2 | 8.4 | 6.5 | 8.2 | 7.5 | 6.8 | 6 | 5 | 7.3 | 6 | 8.2334 | 7.4 | 7.6 | 7 | 7.3 | 7 | 8.5 | 8 | 6 |  | 100
| 1 | 6 | 8.0741 | 7.2 | 6 | 7.3 | 8 | 6.3 | 8.2 | 7.5 | 6.7 | 6 | 5 | 7.5 | 6 | 7.8814 | 6.8 | 7 | 6 | 7.5 | 7 | 8.7 | 10 | 7
| Mean |  | 7.9408 | 7.3167 | 6 | 7.1167 | 8.2833 | 6.4333 | 8.2167 | 7.4667 | 6.7833 | 6 | 5 | 7.3833 | 6 | 8.0497 | 6.75 | 7.4 | 6.5 | 7.3 | 7 | 8.5667 | 8.8333 | 6.1667
| Median |  | 7.9797 | 7.3 | 6 | 7.1 | 8.25 | 6.45 | 8.2 | 7.45 | 6.8 | 6 | 5 | 7.3 | 6 | 8.0511 | 6.6 | 7.55 | 6 | 7.3 | 7 | 8.5 | 9 | 6
| Std Dev. |  | 0.1575 | 0.075277 | 0 | 0.1169 | 0.20412 | 0.21602 | 0.19408 | 0.13663 | 0.040825 | 0 | 0 | 0.18348 | 0 | 0.15649 | 0.33317 | 0.31623 | 0.83666 | 0.14142 | 0 | 0.15055 | 0.75277 | 0.40825
| Rel.Std.Dev. |  | 1.983% | 1.029% | 0% | 1.643% | 2.464% | 3.358% | 2.362% | 1.83% | 0.6018% | 0% | 0% | 2.485% | 0% | 1.944% | 4.936% | 4.273% | 12.87% | 1.937% | 0% | 1.757% | 8.522% | 6.62%
| PDM3 |  | 12.13% | 3.316% | -15.28% | 0.4918% | 16.97% | -9.157% | 16.02% | 5.434% | -4.215% | -15.28% | -29.4% | 4.257% | -15.28% | 13.67% | -4.686% | 4.493% | -8.216% | 3.081% | -1.156% | 20.97% | 24.73% | -12.92%
| Z-Score (Absolute) |  | 0.66 | 0.1 | 1.08 | 0.08 | 0.96 | 0.69 | 0.9 | 0.23 | 0.38 | 1.08 | 1.98 | 0.16 | 1.08 | 0.75 | 0.41 | 0.17 | 0.63 | 0.08 | 0.19 | 1.22 | 1.45 | 0.93
| Analytical results for Se in OREAS 501b (Certified Value 2.78 ppm) | [24 empty] | Certified
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 22 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  | 2
2
| 1 | 1 | 2.3239 | 3 | 3 | 3.2 | 3.2 | 2.5 | 2 | 2 | < 5 | < 10 | 3.1 | 2.3815 | 3 | 3.3 | < 5 | 3 | < 1 | 2 | 2 | < 12 | < 5 |  |  | 1
| 1 | 2 | 2.4429 | 3 | 3 | 3.1 | 3.4 | 2.2 | 3 | 2 | < 5 | < 10 | 3.1 | 2.7438 | 3 | 3.5 | < 5 | 3.7 | < 1 | 2 | 3 | < 12 | < 5 |  |  | 5
| 1 | 3 | 2.4368 | 3 | 3 | 2.9 | 3.1 | 2.3 | 3 | 3 | < 5 | < 10 | 2.7 | 2.5417 | 3 | 3.3 | < 5 | 3.3 | < 1 | 2 | 2 | < 12 | < 5 |  |  | 16
| 1 | 4 | 2.3335 | 2 | 3 | 3 | 3 | 2.3 | 3 | 2 | < 5 | < 10 | 3.2 | 2.8223 | 3 | 3.4 | < 5 | 3.7 | < 1 | 2 | 3 | < 12 | < 5 |  |  | 2.7835
| 1 | 5 | 2.3686 | 2 | 3 | 2.9 | 3.3 | 2.6 | 3 | 3 | < 5 | < 10 | 3.5 | 2.7427 | 3.1 | 3.3 | < 5 | 3.4 | < 1 | 2 | 2 | < 12 | < 5 |  |  | 101
| 1 | 6 | 2.4638 | 3 | 2 | 3.1 | 3.2 | 2.2 | 3 | 2 | < 5 | < 10 | 2.9 | 2.594 | 3 | 3.3 | < 5 | 3.2 | < 1 | 3 | 2 | < 12 | < 5
| Mean |  | 2.3949 | 2.6667 | 2.8333 | 3.0333 | 3.2 | 2.35 | 2.8333 | 2.3333 |  |  | 3.0833 | 2.6377 | 3.0167 | 3.35 |  | 3.3833 |  | 2.1667 | 2.3333
| Median |  | 2.4027 | 3 | 3 | 3.05 | 3.2 | 2.3 | 3 | 2 |  |  | 3.1 | 2.6684 | 3 | 3.3 |  | 3.35 |  | 2 | 2
| Std Dev. |  | 0.060514 | 0.5164 | 0.40825 | 0.12111 | 0.14142 | 0.16432 | 0.40825 | 0.5164 |  |  | 0.27142 | 0.16316 | 0.040825 | 0.083666 |  | 0.27869 |  | 0.40825 | 0.5164
| Rel.Std.Dev. |  | 2.527% | 19.36% | 14.41% | 3.993% | 4.419% | 6.992% | 14.41% | 22.13% |  |  | 8.803% | 6.186% | 1.353% | 2.497% |  | 8.237% |  | 18.84% | 22.13%
| PDM3 |  | -13.96% | -4.198% | 1.79% | 8.975% | 14.96% | -15.57% | 1.79% | -16.17% |  |  | 10.77% | -5.239% | 8.377% | 20.35% |  | 21.55% |  | -22.16% | -16.17%
| Z-Score (Absolute) |  | 0.67 | 0.38 | 0.2 | 0.02 | 0.2 | 0.72 | 0.2 | 0.74 | 2.15 | 7.57 | 0.07 | 0.41 | 0 | 0.36 | 2.15 | 0.4 | 2.19 | 0.92 | 0.74 | 9.74 | 2.15
| Analytical results for Si in OREAS 501b (Indicative Value 0.038 wt.%) | [24 empty] | Indicative
Si | Round | Replicate | 01: Lab |  |  | 1
| No | No | 01: 19 |  |  | wt.%
|  |  | 01: AR*OES |  |  | 3
3
| 1 | 1 | 01: 0.03 |  |  | 1
| 1 | 2 | 01: 0.04 |  |  | 3
| 1 | 3 | 01: 0.05 |  |  | 16
| 1 | 4 | 01: 0.05 |  |  | 0.038333
| 1 | 5 | 01: 0.03 |  |  | 14
| 1 | 6 | 01: 0.03
| Mean |  | 01: 0.038333
| Median |  | 01: 0.035
| Std Dev. |  | 01: 0.0098319
| Rel.Std.Dev. |  | 01: 25.65%
| PDM3 |  | 01: 8.882e-14%
Z-Score (Absolute) | bar: 01
| Analytical results for Sm in OREAS 501b (Indicative Value 4.71 ppm) | [24 empty] | Indicative
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS |  |  | 2
2
| 1 | 1 | 01: < 10; 02: 4.89; 03: 4.43; 04: 4.45; 05: 4.6 |  |  | 1
| 1 | 2 | 01: < 10; 02: 4.62; 03: 4.48; 04: 4.96; 05: 4.8 |  |  | 9
| 1 | 3 | 01: < 10; 02: 4.83; 03: 4.65; 04: 4.56; 05: 4.8 |  |  | 16
| 1 | 4 | 01: < 10; 02: 4.77; 03: 4.55; 04: 4.8; 05: 4.7 |  |  | 4.71
| 1 | 5 | 01: < 10; 02: 4.59; 03: 4.4; 04: 4.72; 05: 4.8 |  |  | 15
| 1 | 6 | 01: < 10; 02: 4.51; 03: 4.36; 04: 4.68; 05: 4.7
| Mean |  | 02: 4.7017; 03: 4.4783; 04: 4.695; 05: 4.7333
| Median |  | 02: 4.695; 03: 4.455; 04: 4.7; 05: 4.75
| Std Dev. |  | 02: 0.14999; 03: 0.10685; 04: 0.17908; 05: 0.08165
| Rel.Std.Dev. |  | 02: 3.19%; 03: 2.386%; 04: 3.814%; 05: 1.725%
| PDM3 |  | 02: -0.1769%; 03: -4.919%; 04: -0.3185%; 05: 0.4954%
| Z-Score (Absolute) |  | 01: 112.82; 02: 0; 03: 4.76; 04: 0.14; 05: 0.67
| Analytical results for Sn in OREAS 501b (Certified Value 4.65 ppm) | [24 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 2
3
| 1 | 1 | 4.3189 | 4.91 | 4.3 | 4.4 | 4.6 | 4.2 | 4.61 | < 10 | < 5 | 4.3 | 5.0641 | 5.2 | 5 | < 10 | 4.1 | < 5 | 4.4 | 4.9 | < 5 | < 5 |  |  |  | 1
| 1 | 2 | 4.3654 | 4.74 | 4.2 | 4.4 | 4.7 | 4.4 | 5.01 | < 10 | < 5 | 4.2 | 5.1027 | 5.2 | 5.1 | < 10 | 4.3 | < 5 | 4.5 | 4.7 | < 5 | < 5 |  |  |  | #N/A
| 1 | 3 | 4.3924 | 4.89 | 4.5 | 4.4 | 4.8 | 4.3 | 5.2 | < 10 | < 5 | 4.3 | 5.2656 | 5.2 | 4.9 | < 10 | 4.3 | < 5 | 4.5 | 5 | < 5 | < 5 |  |  |  | 16
| 1 | 4 | 4.2467 | 4.84 | 4.5 | 4.3 | 4.7 | 4.6 | 4.91 | < 10 | < 5 | 4.2 | 5.1621 | 5.3 | 5 | < 10 | 4.3 | < 5 | 4.4 | 4.7 | < 5 | < 5 |  |  |  | 4.6506
| 1 | 5 | 4.2511 | 4.85 | 4.3 | 4.6 | 4.7 | 4.5 | 4.84 | < 10 | < 5 | 4.3 | 5.2002 | 5.1 | 5 | < 10 | 4.3 | < 5 | 4.3 | 4.9 | < 5 | < 5 |  |  |  | 102
| 1 | 6 | 4.4658 | 4.75 | 4.3 | 4.4 | 4.7 | 4.5 | 4.86 | < 10 | < 5 | 4.4 | 5.0053 | 5.3 | 4.6 | < 10 | 4.4 | < 5 | 4.5 | 4.6 | < 5 | < 5
| Mean |  | 4.34 | 4.83 | 4.35 | 4.4167 | 4.7 | 4.4167 | 4.905 |  |  | 4.2833 | 5.1334 | 5.2167 | 4.9333 |  | 4.2833 |  | 4.4333 | 4.8
| Median |  | 4.3422 | 4.845 | 4.3 | 4.4 | 4.7 | 4.45 | 4.885 |  |  | 4.3 | 5.1324 | 5.2 | 5 |  | 4.3 |  | 4.45 | 4.8
| Std Dev. |  | 0.08517 | 0.070711 | 0.12247 | 0.098319 | 0.063246 | 0.1472 | 0.19563 |  |  | 0.075277 | 0.094774 | 0.075277 | 0.17512 |  | 0.098319 |  | 0.08165 | 0.15492
| Rel.Std.Dev. |  | 1.962% | 1.464% | 2.816% | 2.226% | 1.346% | 3.333% | 3.988% |  |  | 1.757% | 1.846% | 1.443% | 3.55% |  | 2.295% |  | 1.842% | 3.227%
| PDM3 |  | -6.678% | 3.858% | -6.464% | -5.03% | 1.062% | -5.03% | 5.47% |  |  | -7.897% | 10.38% | 12.17% | 6.079% |  | -7.897% |  | -4.672% | 3.212%
| Z-Score (Absolute) |  | 1.16 | 0.08 | 1.13 | 0.99 | 0.37 | 0.99 | 0.08 | 11.25 | 0.29 | 1.28 | 0.58 | 0.77 | 0.14 | 11.25 | 1.28 | 0.29 | 0.95 | 0.15 | 0.29 | 0.29
| Analytical results for Sr in OREAS 501b (Certified Value 60 ppm) | [24 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 0
1
| 1 | 1 | 63.797 | 67.99 | 54.7 | 56.3 | 59.9 | 56 | 53.6 | 60.52 | 103 | 56 | 53 | 62 | 60 | 65.2228 | 56.8 | 65.3 | 34 | 58.6 | 60.7 | 59.9 | 70 | 103 | 85 | 1
| 1 | 2 | 64.141 | 68.77 | 53.2 | 57.9 | 60.7 | 54 | 54 | 63.1 | 102 | 56 | 54 | 59 | 59 | 66.4149 | 58.1 | 66.6 | 36 | 60.9 | 61.2 | 60.1 | 67.5 | 122 | 84 | #N/A
| 1 | 3 | 61.798 | 64.2 | 56.3 | 56.7 | 61.3 | 52 | 53.4 | 61.99 | 103 | 59 | 53 | 60.7 | 59 | 63.7469 | 59.7 | 62.9 | 39 | 60.5 | 63.1 | 61.7 | 67.5 | 108 | 84 | 16
| 1 | 4 | 61.507 | 66.68 | 57.3 | 55.2 | 60 | 52 | 54.1 | 61.88 | 103 | 59 | 55 | 60.5 | 58 | 65.626 | 59.1 | 64.6 | 47 | 61.5 | 62.6 | 62.8 | 67.3 | 127 | 86 | 60.1103
| 1 | 5 | 60.405 | 65.06 | 52.9 | 56.2 | 61.8 | 50 | 53.4 | 64.69 | 103 | 61 | 52 | 61.8 | 58 | 65.5519 | 63.8 | 64.4 | 40 | 60.4 | 61.7 | 59.8 | 68.8 | 87 | 84 | 103
| 1 | 6 | 63.312 | 64.37 | 55 | 57.2 | 59.9 | 51 | 54.1 | 63.58 | 100 | 60 | 53 | 62.3 | 58 | 62.5235 | 59.9 | 59.6 | 37 | 62 | 62.5 | 64.5 | 67.1 | 114 | 87
| Mean |  | 62.4933 | 66.1783 | 54.9 | 56.5833 | 60.6 | 52.5 | 53.7667 | 62.6267 | 102.3333 | 58.5 | 53.3333 | 61.05 | 58.6667 | 64.8477 | 59.5667 | 63.9 | 38.8333 | 60.65 | 61.9667 | 61.4667 | 68.0333 | 110.1667 | 85
| Median |  | 62.555 | 65.87 | 54.85 | 56.5 | 60.35 | 52 | 53.8 | 62.545 | 103 | 59 | 53 | 61.25 | 58.5 | 65.3873 | 59.4 | 64.5 | 38 | 60.7 | 62.1 | 60.9 | 67.5 | 111 | 84.5
| Std Dev. |  | 1.4766 | 1.9331 | 1.7123 | 0.92394 | 0.80994 | 2.1679 | 0.33862 | 1.4691 | 1.2111 | 2.0736 | 1.0328 | 1.2373 | 0.8165 | 1.4361 | 2.3695 | 2.4282 | 4.535 | 1.1743 | 0.92014 | 1.9044 | 1.1343 | 14.3585 | 1.2649
| Rel.Std.Dev. |  | 2.363% | 2.921% | 3.119% | 1.633% | 1.337% | 4.129% | 0.6298% | 2.346% | 1.183% | 3.545% | 1.936% | 2.027% | 1.392% | 2.215% | 3.978% | 3.8% | 11.68% | 1.936% | 1.485% | 3.098% | 1.667% | 13.03% | 1.488%
| PDM3 |  | 3.964% | 10.09% | -8.668% | -5.868% | 0.8146% | -12.66% | -10.55% | 4.186% | 70.24% | -2.679% | -11.27% | 1.563% | -2.402% | 7.881% | -0.9044% | 6.305% | -35.4% | 0.8978% | 3.088% | 2.256% | 13.18% | 83.27% | 41.41%
| Z-Score (Absolute) |  | 0.26 | 0.91 | 1.09 | 0.79 | 0.08 | 1.52 | 1.29 | 0.28 | 7.33 | 0.45 | 1.37 | 0 | 0.42 | 0.67 | 0.26 | 0.51 | 3.94 | 0.07 | 0.16 | 0.07 | 1.24 | 8.72 | 4.25
| Analytical results for Ta in OREAS 501b (Indicative Value 0.7 ppm) | [24 empty] | Indicative
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 07 | 08 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 27 | [7 empty] | ppm
|  |  | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES | [7 empty] | 2
2
| 1 | 1 | < 0.01 | < 0.005 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.052174 | 0.01 | 0.01 | 4 | 0.01 | < 0.1 | < 0.05 | < 5 | [7 empty] | 1
| 1 | 2 | < 0.01 | < 0.005 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.038382 | 0.01 | 0.01 | 5 | 0.01 | < 0.1 | < 0.05 | < 5 | [7 empty] | 10
| 1 | 3 | < 0.01 | < 0.005 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.028125 | 0.01 | 0.01 | 5 | 0.01 | < 0.1 | < 0.05 | < 5 | [7 empty] | 16
| 1 | 4 | < 0.01 | < 0.005 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.021484 | 0.01 | 0.01 | 9 | 0.01 | < 0.1 | < 0.05 | < 5 | [7 empty] | 0.69643
| 1 | 5 | < 0.01 | < 0.005 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.020747 | 0.01 | 0.01 | 5 | 0.01 | < 0.1 | < 0.05 | < 5 | [7 empty] | 16
| 1 | 6 | < 0.01 | 0.005 | 0.01 | < 0.01 | < 0.05 | < 0.01 | < 5 | < 0.01 | 0.019231 | 0.01 | 0.01 | 5 | 0.01 | < 0.1 | < 0.05 | < 5
| Mean |  | 02: 0.005; 03: 0.01; 09: 0.030024; 10: 0.01; 11: 0.01; 12: 5.5; 15: 0.01
| Median |  | 02: 0.005; 03: 0.01; 09: 0.024805; 10: 0.01; 11: 0.01; 12: 5; 15: 0.01
| Std Dev. |  | 03: 0; 09: 0.012958; 10: 0; 11: 0; 12: 1.7607; 15: 0
| Rel.Std.Dev. |  | 03: 0%; 09: 43.16%; 10: 0%; 11: 0%; 12: 32.01%; 15: 0%
| PDM3 |  | 02: -99.28%; 03: -98.56%; 09: -95.69%; 10: -98.56%; 11: -98.56%; 12: 689.7%; 15: -98.56%
| Z-Score (Absolute) |  | 0 | 1.35 | 0 | 0 | 10.79 | 0 | 1345.92 | 0 | 5.4 | 0 | 0 | 1480.78 | 0 | 24.28 | 10.79 | 1345.92
| Analytical results for Tb in OREAS 501b (Certified Value 0.56 ppm) | [24 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 07; 05: 08; 06: 16; 07: 23; 08: 25 | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*OES/MS; 05: AR*MS; 06: AR*OES/MS; 07: AR*OES/MS; 08: AR*OES/MS | 2
3
| 1 | 1 | 01: < 5; 02: 0.587; 03: 0.515; 04: 0.6; 05: 0.562; 06: 0.571; 07: 0.5; 08: 0.52 | 1
| 1 | 2 | 01: < 5; 02: 0.549; 03: 0.52; 04: 0.61; 05: 0.595; 06: 0.579; 07: 0.5; 08: 0.5 | #N/A
| 1 | 3 | 01: < 5; 02: 0.559; 03: 0.535; 04: 0.59; 05: 0.562; 06: 0.593; 07: 0.5; 08: 0.51 | 16
| 1 | 4 | 01: < 5; 02: 0.559; 03: 0.53; 04: 0.61; 05: 0.587; 06: 0.599; 07: 0.6; 08: 0.51 | 0.55787
| 1 | 5 | 01: < 5; 02: 0.558; 03: 0.505; 04: 0.6; 05: 0.597; 06: 0.591; 07: 0.6; 08: 0.5 | 104
| 1 | 6 | 01: < 5; 02: 0.553; 03: 0.505; 04: 0.6; 05: 0.595; 06: 0.576; 07: 0.6; 08: 0.53
| Mean |  | 02: 0.56083; 03: 0.51833; 04: 0.60167; 05: 0.583; 06: 0.58483; 07: 0.55; 08: 0.51167
| Median |  | 02: 0.5585; 03: 0.5175; 04: 0.6; 05: 0.591; 06: 0.585; 07: 0.55; 08: 0.51
| Std Dev. |  | 02: 0.013423; 03: 0.012517; 04: 0.0075277; 05: 0.016625; 06: 0.011035; 07: 0.054772; 08: 0.01169
| Rel.Std.Dev. |  | 02: 2.393%; 03: 2.415%; 04: 1.251%; 05: 2.852%; 06: 1.887%; 07: 9.959%; 08: 2.285%
| PDM3 |  | 02: 0.5309%; 03: -7.087%; 04: 7.85%; 05: 4.504%; 06: 4.833%; 07: -1.411%; 08: -8.282%
| Z-Score (Absolute) |  | 01: 115.58; 02: 0.29; 03: 1.4; 04: 0.78; 05: 0.29; 06: 0.34; 07: 0.57; 08: 1.57
| Analytical results for Te in OREAS 501b (Certified Value 0.078 ppm) | [24 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 3
3
| 1 | 1 | 0.107 | 0.05 | 0.08 | 0.06 | 0.1 | < 0.2 | 0.09 | 0.1 | < 5 | 0.09 | 0.10435 | 0.09 | 0.06 | < 1 | 0.09 | < 1 | 0.07 | < 0.05 | < 1 | 9 |  |  |  | 1
| 1 | 2 | 0.1593 | 0.06 | 0.08 | 0.07 | 0.12 | < 0.2 | 0.08 | 0.07 | < 5 | 0.06 | 0.0861 | 0.1 | 0.07 | < 1 | 0.06 | < 1 | 0.06 | < 0.05 | < 1 | 9 |  |  |  | #N/A
| 1 | 3 | 0.1149 | 0.08 | 0.08 | 0.07 | 0.08 | < 0.2 | 0.06 | 0.09 | < 5 | 0.08 | 0.090625 | 0.09 | 0.06 | < 1 | 0.1 | < 1 | 0.07 | < 0.05 | < 1 | 9 |  |  |  | 16
| 1 | 4 | 0.1267 | 0.05 | 0.08 | 0.07 | 0.1 | < 0.2 | 0.07 | 0.08 | < 5 | 0.08 | 0.076172 | 0.09 | 0.07 | < 1 | 0.07 | < 1 | 0.09 | < 0.05 | < 1 | 9 |  |  |  | 0.077994
| 1 | 5 | 0.1197 | 0.02 | 0.08 | 0.08 | 0.09 | < 0.2 | 0.08 | 0.07 | < 5 | 0.08 | 0.087137 | 0.09 | 0.07 | < 1 | 0.08 | < 1 | 0.07 | < 0.05 | < 1 | 9 |  |  |  | 105
| 1 | 6 | 0.1304 | 0.02 | 0.1 | 0.07 | 0.09 | < 0.2 | 0.06 | 0.1 | < 5 | 0.06 | 0.097222 | 0.11 | 0.06 | 2 | 0.07 | < 1 | 0.05 | < 0.05 | < 1 | 10
| Mean |  | 0.12633 | 0.046667 | 0.083333 | 0.07 | 0.096667 |  | 0.073333 | 0.085 |  | 0.075 | 0.090267 | 0.095 | 0.065 | 2 | 0.078333 |  | 0.068333 |  |  | 9.1667
| Median |  | 0.1232 | 0.05 | 0.08 | 0.07 | 0.095 |  | 0.075 | 0.085 |  | 0.08 | 0.088881 | 0.09 | 0.065 | 2 | 0.075 |  | 0.07 |  |  | 9
| Std Dev. |  | 0.018177 | 0.023381 | 0.008165 | 0.0063246 | 0.013663 |  | 0.012111 | 0.013784 |  | 0.012247 | 0.0097222 | 0.0083666 | 0.0054772 |  | 0.01472 |  | 0.013292 |  |  | 0.40825
| Rel.Std.Dev. |  | 14.39% | 50.1% | 9.798% | 9.035% | 14.13% |  | 16.51% | 16.22% |  | 16.33% | 10.77% | 8.807% | 8.427% |  | 18.79% |  | 19.45% |  |  | 4.454%
| PDM3 |  | 61.98% | -40.17% | 6.845% | -10.25% | 23.94% |  | -5.976% | 8.982% |  | -3.839% | 15.74% | 21.8% | -16.66% | 2464% | 0.4344% |  | -12.39% |  |  | 11653%
| Z-Score (Absolute) |  | 1.24 | 1.32 | 0.14 | 0.57 | 0.29 | 3.61 | 0.46 | 0.08 | 157.98 | 0.41 | 0.08 | 0.24 | 0.73 | 34.7 | 0.3 | 29.34 | 0.62 | 1.21 | 29.34 | 291.99
| Analytical results for Th in OREAS 501b (Certified Value 17.9 ppm) | [24 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [5 empty] | 1
2
| 1 | 1 | 17.833 | 17.6 | 18.8 | 16.5 | 17.8 | 18.7 | 19 | 16.7 | 18 | 18.4 | 20 | 18.8 | 16.5 | 18 | < 50 | 17.1 | 0.51 | 15.4 | [5 empty] | 1
| 1 | 2 | 19.037 | 16.93 | 18.6 | 16.5 | 18.5 | 18.8 | 18.1 | 17.97 | 17 | 17.7 | 20 | 17.22 | 15.9 | 18.6 | < 50 | 18.7 | 0.52 | 15.6 | [5 empty] | #N/A
| 1 | 3 | 18.768 | 17.8 | 18.8 | 16.2 | 17.8 | 17.7 | 17.2 | 17.81 | 19 | 15.5 | 20 | 17.53 | 16.3 | 18.7 | < 50 | 17.8 | 0.5 | 15.4 | [5 empty] | 16
| 1 | 4 | 18.541 | 16.86 | 19.7 | 17.2 | 18.8 | 17.7 | 18.8 | 18.9 | 18 | 19.6 | 20 | 17.62 | 18.7 | 20.9 | < 50 | 18.3 | 0.54 | 16.1 | [5 empty] | 17.9353
| 1 | 5 | 18.171 | 15.53 | 18.6 | 17.5 | 19.9 | 18.6 | 17.2 | 18.01 | 17 | 18.2 | 20 | 16.38 | 19.6 | 19 | < 50 | 19.2 | 0.52 | 16.1 | [5 empty] | 106
| 1 | 6 | 19.178 | 17.41 | 18.1 | 16.7 | 18.7 | 17.5 | 18.1 | 17.41 | 18 | 17.6 | 20 | 16.58 | 16.9 | 17.4 | < 50 | 18.8 | 0.54 | 16.1
| Mean |  | 18.588 | 17.0217 | 18.7667 | 16.7667 | 18.5833 | 18.1667 | 18.0667 | 17.8 | 17.8333 | 17.8333 | 20 | 17.355 | 17.3167 | 18.7667 |  | 18.3167 | 0.52167 | 15.7833
| Median |  | 18.6545 | 17.17 | 18.7 | 16.6 | 18.6 | 18.15 | 18.1 | 17.89 | 18 | 17.95 | 20 | 17.375 | 16.7 | 18.65 |  | 18.5 | 0.52 | 15.85
| Std Dev. |  | 0.51541 | 0.81862 | 0.52409 | 0.48854 | 0.77825 | 0.59217 | 0.76333 | 0.72735 | 0.75277 | 1.3486 | 0 | 0.86699 | 1.4838 | 1.1911 |  | 0.76267 | 0.016021 | 0.35449
| Rel.Std.Dev. |  | 2.773% | 4.809% | 2.793% | 2.914% | 4.188% | 3.26% | 4.225% | 4.086% | 4.221% | 7.562% | 0% | 4.996% | 8.569% | 6.347% |  | 4.164% | 3.071% | 2.246%
| PDM3 |  | 3.639% | -5.094% | 4.635% | -6.516% | 3.613% | 1.29% | 0.7325% | -0.7543% | -0.5685% | -0.5685% | 11.51% | -3.235% | -3.449% | 4.635% |  | 2.126% | -97.09% | -12%
| Z-Score (Absolute) |  | 0.68 | 0.98 | 0.87 | 1.26 | 0.67 | 0.23 | 0.12 | 0.16 | 0.12 | 0.12 | 2.17 | 0.63 | 0.67 | 0.87 | 34 | 0.39 | 18.49 | 2.3
| Analytical results for Ti in OREAS 501b (Certified Value 0.338 wt.%) | [24 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 3
3
| 1 | 1 | 0.37877 | > 0 | 0.29 | 0.327 | 0.34 | 0.32 | 0.34 | 0.3657 | 0.04 | 0.29 | 0.313 | 0.352 | 0.32 | 0.36113 | 0.34 | 0.348 | 0.2685 | 0.339 | 0.344 | 0.36 | 0.37 | 0.34 | 0.31 | 1
| 1 | 2 | 0.38704 | > 0 | 0.27 | 0.338 | 0.337 | 0.303 | 0.34 | 0.3691 | 0.04 | 0.28 | 0.315 | 0.35 | 0.34 | 0.36967 | 0.344 | 0.338 | 0.2714 | 0.364 | 0.353 | 0.36 | 0.37 | 0.31 | 0.31 | #N/A
| 1 | 3 | 0.37097 | > 0 | 0.295 | 0.326 | 0.342 | 0.296 | 0.34 | 0.3551 | 0.04 | 0.29 | 0.319 | 0.351 | 0.33 | 0.35982 | 0.349 | 0.347 | 0.3011 | 0.353 | 0.351 | 0.38 | 0.36 | 0.35 | 0.32 | 16
| 1 | 4 | 0.37693 | > 0 | 0.295 | 0.332 | 0.343 | 0.302 | 0.35 | 0.3601 | 0.04 | 0.29 | 0.323 | 0.357 | 0.32 | 0.3722 | 0.351 | 0.351 | 0.3304 | 0.349 | 0.34 | 0.39 | 0.35 | 0.37 | 0.32 | 0.33778
| 1 | 5 | 0.36856 | > 0 | 0.285 | 0.335 | 0.337 | 0.29 | 0.34 | 0.3625 | 0.04 | 0.29 | 0.32 | 0.357 | 0.32 | 0.36694 | 0.359 | 0.339 | 0.3166 | 0.35 | 0.336 | 0.37 | 0.37 | 0.37 | 0.32 | 107
| 1 | 6 | 0.37963 | > 0 | 0.295 | 0.341 | 0.339 | 0.288 | 0.35 | 0.3533 | 0.04 | 0.29 | 0.322 | 0.361 | 0.32 | 0.36518 | 0.361 | 0.338 | 0.2933 | 0.351 | 0.351 | 0.4 | 0.35 | 0.37 | 0.33
| Mean |  | 0.37698 |  | 0.28833 | 0.33317 | 0.33967 | 0.29983 | 0.34333 | 0.36097 | 0.04 | 0.28833 | 0.31867 | 0.35467 | 0.325 | 0.36582 | 0.35067 | 0.3435 | 0.29688 | 0.351 | 0.34583 | 0.37667 | 0.36167 | 0.35167 | 0.31833
| Median |  | 0.37785 |  | 0.2925 | 0.3335 | 0.3395 | 0.299 | 0.34 | 0.3613 | 0.04 | 0.29 | 0.3195 | 0.3545 | 0.32 | 0.36606 | 0.35 | 0.343 | 0.2972 | 0.3505 | 0.3475 | 0.375 | 0.365 | 0.36 | 0.32
| Std Dev. |  | 0.0066098 |  | 0.0098319 | 0.0059805 | 0.0025033 | 0.0116 | 0.005164 | 0.0060804 | 0 | 0.0040825 | 0.0039328 | 0.0043205 | 0.0083666 | 0.0047978 | 0.0082138 | 0.0058224 | 0.024486 | 0.008025 | 0.0069113 | 0.01633 | 0.0098319 | 0.024014 | 0.0075277
| Rel.Std.Dev. |  | 1.753% |  | 3.41% | 1.795% | 0.737% | 3.869% | 1.504% | 1.684% | 0% | 1.416% | 1.234% | 1.218% | 2.574% | 1.312% | 2.342% | 1.695% | 8.248% | 2.286% | 1.998% | 4.335% | 2.719% | 6.829% | 2.365%
| PDM3 |  | 11.61% |  | -14.64% | -1.365% | 0.5594% | -11.23% | 1.645% | 6.865% | -88.16% | -14.64% | -5.658% | 5% | -3.783% | 8.303% | 3.816% | 1.694% | -12.11% | 3.915% | 2.385% | 11.51% | 7.073% | 4.112% | -5.756%
| Z-Score (Absolute) |  | 1.23 | NA | 2.03 | 0.38 | 0.14 | 1.6 | 0 | 0.65 | 11.16 | 2.03 | 0.91 | 0.41 | 0.68 | 0.82 | 0.27 | 0 | 1.71 | 0.28 | 0.09 | 1.22 | 0.67 | 0.3 | 0.92
| Analytical results for Tl in OREAS 501b (Certified Value 0.69 ppm) | [24 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 27 |  |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  |  |  | 2
3
| 1 | 1 | 0.6837 | 0.72 | 0.66 | 0.69 | 0.69 | 0.7 | 0.7 | 0.69 | < 5 | 0.65 | < 10 | 0.78913 | 0.66 | 0.72 | 5 | 0.6 | 0.84 | 0.56 | < 10 |  |  |  |  | 1
| 1 | 2 | 0.7166 | 0.7 | 0.66 | 0.71 | 0.7 | 0.6 | 0.72 | 0.72 | < 5 | 0.66 | < 10 | 0.77386 | 0.65 | 0.72 | < 2 | 0.64 | 0.83 | 0.55 | < 10 |  |  |  |  | #N/A
| 1 | 3 | 0.6727 | 0.72 | 0.69 | 0.68 | 0.71 | 0.6 | 0.72 | 0.7 | < 5 | 0.64 | < 10 | 0.79896 | 0.66 | 0.72 | 3 | 0.62 | 0.89 | 0.55 | < 10 |  |  |  |  | 16
| 1 | 4 | 0.6665 | 0.7 | 0.69 | 0.69 | 0.72 | 0.7 | 0.73 | 0.7 | < 5 | 0.67 | < 10 | 0.78125 | 0.65 | 0.72 | 5 | 0.61 | 0.91 | 0.57 | < 10 |  |  |  |  | 0.69367
| 1 | 5 | 0.6716 | 0.71 | 0.65 | 0.69 | 0.7 | 0.8 | 0.72 | 0.72 | < 5 | 0.67 | < 10 | 0.78527 | 0.69 | 0.71 | < 2 | 0.61 | 0.89 | 0.57 | < 10 |  |  |  |  | 108
| 1 | 6 | 0.6962 | 0.71 | 0.65 | 0.7 | 0.71 | 0.7 | 0.73 | 0.71 | < 5 | 0.68 | < 10 | 0.77244 | 0.69 | 0.64 | < 2 | 0.63 | 0.83 | 0.58 | < 10
| Mean |  | 0.68455 | 0.71 | 0.66667 | 0.69333 | 0.705 | 0.68333 | 0.72 | 0.70667 |  | 0.66167 |  | 0.78348 | 0.66667 | 0.705 | 4.3333 | 0.61833 | 0.865 | 0.56333
| Median |  | 0.6782 | 0.71 | 0.66 | 0.69 | 0.705 | 0.7 | 0.72 | 0.705 |  | 0.665 |  | 0.78326 | 0.66 | 0.72 | 5 | 0.615 | 0.865 | 0.565
| Std Dev. |  | 0.018961 | 0.0089443 | 0.018619 | 0.010328 | 0.010488 | 0.075277 | 0.010954 | 0.012111 |  | 0.01472 |  | 0.009942 | 0.018619 | 0.032094 | 1.1547 | 0.01472 | 0.035637 | 0.012111
| Rel.Std.Dev. |  | 2.77% | 1.26% | 2.793% | 1.49% | 1.488% | 11.02% | 1.521% | 1.714% |  | 2.225% |  | 1.269% | 2.793% | 4.552% | 26.65% | 2.381% | 4.12% | 2.15%
| PDM3 |  | -1.315% | 2.354% | -3.893% | -0.04843% | 1.633% | -1.49% | 3.796% | 1.874% |  | -4.614% |  | 12.95% | -3.893% | 1.633% | 524.7% | -10.86% | 24.7% | -18.79%
| Z-Score (Absolute) |  | 0.36 | 0.09 | 0.67 | 0.21 | 0 | 0.38 | 0.26 | 0.03 | 75.55 | 0.76 | 163.51 | 1.38 | 0.67 | 0 | 43.3 | 1.52 | 2.81 | 2.49 | 163.51
| Analytical results for Tm in OREAS 501b (Indicative Value 0.23 ppm) | [24 empty] | Indicative
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab |  |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23 |  |  |  |  | ppm
|  |  | 01: AR*MS; 02: AR*OES/MS; 03: AR*MS; 04: AR*MS; 05: AR*OES/MS |  |  |  |  | 2
2
| 1 | 1 | 01: < 5; 02: 0.26; 03: 0.215; 04: 0.25; 05: 0.2 |  |  |  |  | 1
| 1 | 2 | 01: < 5; 02: 0.25; 03: 0.21; 04: 0.26; 05: 0.2 |  |  |  |  | 11
| 1 | 3 | 01: < 5; 02: 0.25; 03: 0.22; 04: 0.25; 05: 0.2 |  |  |  |  | 16
| 1 | 4 | 01: < 5; 02: 0.24; 03: 0.225; 04: 0.26; 05: 0.2 |  |  |  |  | 0.22896
| 1 | 5 | 01: < 5; 02: 0.25; 03: 0.215; 04: 0.25; 05: 0.2 |  |  |  |  | 17
| 1 | 6 | 01: < 5; 02: 0.24; 03: 0.21; 04: 0.24; 05: 0.2
| Mean |  | 02: 0.24833; 03: 0.21583; 04: 0.25167; 05: 0.2
| Median |  | 02: 0.25; 03: 0.215; 04: 0.25; 05: 0.2
| Std Dev. |  | 02: 0.0075277; 03: 0.0058452; 04: 0.0075277; 05: 3.0405e-17
| Rel.Std.Dev. |  | 02: 3.031%; 03: 2.708%; 04: 2.991%; 05: 1.52e-14%
| PDM3 |  | 02: 8.462%; 03: -5.732%; 04: 9.918%; 05: -12.65%
| Z-Score (Absolute) |  | 01: 98.59; 02: 0; 03: 0.67; 04: 0.07; 05: 1
| Analytical results for U in OREAS 501b (Certified Value 4.56 ppm) | [24 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | [7 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | [7 empty] | 2
3
| 1 | 1 | 4.7442 | 4.52 | 4.32 | 4.43 | 4.28 | 4.8 | 4.95 | 4.79 | < 10 | 4.568 | 3.99 | 4.37 | 21 | 4.3 | 4.56 | 4.15 | [7 empty] | 1
| 1 | 2 | 4.8546 | 4.57 | 4.14 | 4.32 | 4.35 | 4.72 | 4.75 | 4.66 | < 10 | 4.764 | 3.76 | 4.74 | 21 | 4.48 | 4.57 | 4.06 | [7 empty] | #N/A
| 1 | 3 | 4.7379 | 4.58 | 4.46 | 4.33 | 4.25 | 4.62 | 5 | 4.11 | < 10 | 4.355 | 4.18 | 4.37 | 23 | 4.7 | 4.63 | 3.69 | [7 empty] | 16
| 1 | 4 | 4.748 | 4.72 | 4.6 | 4.37 | 4.42 | 5.26 | 5.18 | 4.89 | < 10 | 4.792 | 4.22 | 5.52 | 26 | 5.1 | 4.66 | 4.15 | [7 empty] | 4.5575
| 1 | 5 | 4.6686 | 4.68 | 4.36 | 4.45 | 5.2 | 4.49 | 5.2 | 4.39 | < 10 | 4.235 | 4.49 | 4.52 | 20 | 4.67 | 4.72 | 4.19 | [7 empty] | 109
| 1 | 6 | 4.8746 | 4.64 | 4.32 | 4.85 | 4.99 | 4.77 | 4.85 | 4.17 | < 10 | 4.151 | 4.08 | 4.46 | 26 | 4.41 | 4.78 | 4.02
| Mean |  | 4.7713 | 4.6183 | 4.3667 | 4.4583 | 4.5817 | 4.7767 | 4.9883 | 4.5017 |  | 4.4775 | 4.12 | 4.6633 | 22.8333 | 4.61 | 4.6533 | 4.0433
| Median |  | 4.7461 | 4.61 | 4.34 | 4.4 | 4.385 | 4.745 | 4.975 | 4.525 |  | 4.4615 | 4.13 | 4.49 | 22 | 4.575 | 4.645 | 4.105
| Std Dev. |  | 0.078166 | 0.074944 | 0.15423 | 0.19884 | 0.4074 | 0.26235 | 0.17837 | 0.32707 |  | 0.27188 | 0.24438 | 0.4413 | 2.6394 | 0.28453 | 0.085713 | 0.18435
| Rel.Std.Dev. |  | 1.638% | 1.623% | 3.532% | 4.46% | 8.892% | 5.492% | 3.576% | 7.266% |  | 6.072% | 5.931% | 9.463% | 11.56% | 6.172% | 1.842% | 4.559%
| PDM3 |  | 4.692% | 1.335% | -4.187% | -2.176% | 0.5303% | 4.809% | 9.453% | -1.225% |  | -1.755% | -9.6% | 2.322% | 401% | 1.152% | 2.103% | -11.28%
| Z-Score (Absolute) |  | 0.68 | 0.02 | 1.07 | 0.67 | 0.14 | 0.7 | 1.61 | 0.48 | 23.21 | 0.59 | 2.13 | 0.21 | 78.5 | 0.02 | 0.17 | 2.46
| Analytical results for V in OREAS 501b (Certified Value 110 ppm) | [24 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  | 0
0
| 1 | 1 | 120.48 | 116 | 106 | 107 | 108 | 109 | 118 | 59 | 107 | 108 | 112 | 98 | 110.913 | 105 | 110 | 93 | 108 | 115 | 112 | 101 | 108 | 105 |  | 1
| 1 | 2 | 121.74 | 117 | 108 | 106 | 111 | 112 | 121 | 60 | 109 | 109 | 111 | 104 | 114.1494 | 107 | 106 | 94 | 111 | 118 | 112 | 100 | 98 | 105 |  | #N/A
| 1 | 3 | 117 | 116 | 105 | 107 | 109 | 108 | 119 | 59 | 110 | 109 | 111 | 104 | 111.375 | 108 | 109 | 103 | 111 | 114 | 117 | 105 | 103 | 105 |  | 16
| 1 | 4 | 118.27 | 119 | 106 | 108 | 109 | 114 | 119 | 59 | 108 | 110 | 113 | 99 | 114.3555 | 108 | 111 | 122 | 111 | 112 | 119 | 100 | 109 | 105 |  | 109.7369
| 1 | 5 | 116.48 | 118 | 107 | 106 | 109 | 114 | 122 | 60 | 109 | 108 | 113 | 99 | 115.8402 | 110 | 107 | 106 | 110 | 111 | 113 | 99 | 113 | 106 |  | 110
| 1 | 6 | 119 | 117 | 109 | 107 | 109 | 112 | 119 | 58 | 109 | 111 | 114 | 100 | 112.2436 | 111 | 106 | 99 | 111 | 116 | 118 | 99 | 108 | 109
| Mean |  | 118.8283 | 117.1667 | 106.8333 | 106.8333 | 109.1667 | 111.5 | 119.6667 | 59.1667 | 108.6667 | 109.1667 | 112.3333 | 100.6667 | 113.1461 | 108.1667 | 108.1667 | 102.8333 | 110.3333 | 114.3333 | 115.1667 | 100.6667 | 106.5 | 105.8333
| Median |  | 118.635 | 117 | 106.5 | 107 | 109 | 112 | 119 | 59 | 109 | 109 | 112.5 | 99.5 | 113.1965 | 108 | 108 | 101 | 111 | 114.5 | 115 | 100 | 108 | 105
| Std Dev. |  | 2.0201 | 1.169 | 1.472 | 0.75277 | 0.98319 | 2.51 | 1.5055 | 0.75277 | 1.0328 | 1.169 | 1.2111 | 2.6583 | 1.9321 | 2.137 | 2.137 | 10.6474 | 1.2111 | 2.582 | 3.1885 | 2.2509 | 5.244 | 1.6021
| Rel.Std.Dev. |  | 1.7% | 0.9978% | 1.378% | 0.7046% | 0.9006% | 2.251% | 1.258% | 1.272% | 0.9504% | 1.071% | 1.078% | 2.641% | 1.708% | 1.976% | 1.976% | 10.35% | 1.098% | 2.258% | 2.769% | 2.236% | 4.924% | 1.514%
| PDM3 |  | 8.285% | 6.771% | -2.646% | -2.646% | -0.5196% | 1.607% | 9.049% | -46.08% | -0.9753% | -0.5196% | 2.366% | -8.265% | 3.107% | -1.431% | -1.431% | -6.291% | 0.5435% | 4.189% | 4.948% | -8.265% | -2.95% | -3.557%
| Z-Score (Absolute) |  | 2.06 | 1.71 | 0.43 | 0.43 | 0.05 | 0.54 | 2.23 | 10.32 | 0.05 | 0.05 | 0.71 | 1.71 | 0.88 | 0.16 | 0.16 | 1.26 | 0.29 | 1.12 | 1.3 | 1.71 | 0.5 | 0.64
| Analytical results for W in OREAS 501b (Certified Value 2.34 ppm) | [24 empty] | Certified
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 27 |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES |  |  | 2
2
| 1 | 1 | 2.3495 | 1.94 | 1.4 | 2.79 | 1.92 | 1.7 | 2.4 | 2.97 | < 10 | < 10 | < 2 | 1.98 | < 10 | 2.46 | 2.53 | 1.79 | < 10 | 2.04 | 2 | 3.5 | < 5 |  |  | 1
| 1 | 2 | 2.8708 | 2.47 | 1.2 | 2.44 | 1.99 | 1.8 | 2.2 | 2.62 | < 10 | < 10 | 2 | 2.54 | < 10 | 2.86 | 2.52 | 2.09 | < 10 | 2.41 | 2 | 3.6 | < 5 |  |  | 4
| 1 | 3 | 2.2935 | 2.53 | 1.3 | 2.08 | 3.09 | 1.6 | 2.1 | 2.89 | < 10 | < 10 | < 2 | 1.91 | < 10 | 2.87 | 2.45 | 2.79 | < 10 | 3.05 | 4 | 4.8 | < 5 |  |  | 16
| 1 | 4 | 2.4499 | 2.99 | 1.4 | 2.84 | 2.15 | 1.6 | 3.3 | 3.59 | < 10 | < 10 | < 2 | 2.07 | < 10 | 2.96 | 2.58 | 2.27 | < 10 | 2.06 | 2 | 4.2 | < 5 |  |  | 2.3445
| 1 | 5 | 2.3942 | 5.3 | 1.3 | 2.4 | 2.2 | 1.7 | 3.2 | 2.34 | < 10 | < 10 | < 2 | 2.4 | < 10 | 2.46 | 2.42 | 2.74 | < 10 | 1.93 | 2 | 4.2 | < 5 |  |  | 111
| 1 | 6 | 2.3468 | 3.22 | 1.3 | 2.35 | 1.82 | 1.6 | 2 | 2.55 | < 10 | < 10 | 2 | 2.39 | < 10 | 2.71 | 2.64 | 2.34 | < 10 | 2.34 | 3 | 5.1 | < 5
| Mean |  | 2.4508 | 3.075 | 1.3167 | 2.4833 | 2.195 | 1.6667 | 2.5333 | 2.8267 |  |  | 2 | 2.215 |  | 2.72 | 2.5233 | 2.3367 |  | 2.305 | 2.5 | 4.2333
| Median |  | 2.3718 | 2.76 | 1.3 | 2.42 | 2.07 | 1.65 | 2.3 | 2.755 |  |  | 2 | 2.23 |  | 2.785 | 2.525 | 2.305 |  | 2.2 | 2 | 4.2
| Std Dev. |  | 0.21233 | 1.1773 | 0.075277 | 0.28654 | 0.46073 | 0.08165 | 0.57155 | 0.43894 |  |  | 0 | 0.26067 |  | 0.21679 | 0.081158 | 0.38271 |  | 0.40943 | 0.83666 | 0.63456
| Rel.Std.Dev. |  | 8.664% | 38.29% | 5.717% | 11.54% | 20.99% | 4.899% | 22.56% | 15.53% |  |  | 0% | 11.77% |  | 7.97% | 3.216% | 16.38% |  | 17.76% | 33.47% | 14.99%
| PDM3 |  | 4.534% | 31.16% | -43.84% | 5.922% | -6.376% | -28.91% | 8.055% | 20.57% |  |  | -14.69% | -5.523% |  | 16.02% | 7.629% | -0.3334% |  | -1.684% | 6.633% | 80.57%
| Z-Score (Absolute) |  | 0.15 | 1.13 | 2.48 | 0.08 | 0.67 | 1.76 | 0.02 | 0.62 | 15.36 | 15.36 | 1.07 | 0.63 | 15.36 | 0.4 | 0 | 0.38 | 15.36 | 0.45 | 0.05 | 3.51 | 5.09
| Analytical results for Y in OREAS 501b (Certified Value 16.6 ppm) | [24 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 27 | [5 empty] | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*OES | [5 empty] | 1
2
| 1 | 1 | 17.452 | 16.5 | 15.5 | 16 | 16.25 | 16.3 | 16.03 | 7 | 14 | 16.85 | 17.2652 | 14.75 | 17.15 | 13 | 16.2 | 17 | 18.83 | 16 | [5 empty] | 1
| 1 | 2 | 18.199 | 16.63 | 15.1 | 16.35 | 16.65 | 16.9 | 16.68 | 7 | 14 | 16.1 | 17.6473 | 14.95 | 17.2 | 13 | 17.2 | 17 | 18.33 | 16 | [5 empty] | #N/A
| 1 | 3 | 18.047 | 16.99 | 16.1 | 15.8 | 17.05 | 16.8 | 16.4 | 6 | 15 | 16.75 | 17.926 | 15.2 | 16.45 | 15 | 16.65 | 17 | 18.1 | 16 | [5 empty] | 16
| 1 | 4 | 17.423 | 16.44 | 16 | 15.85 | 16.7 | 17.1 | 16.41 | 6 | 14 | 16.3 | 18.0166 | 15.15 | 17.45 | 18 | 17.2 | 16 | 17.76 | 16 | [5 empty] | 16.5936
| 1 | 5 | 17.578 | 16.7 | 15.1 | 15.9 | 16.85 | 16.8 | 16.97 | 7 | 15 | 16.55 | 18.0218 | 16.2 | 17.2 | 15 | 16.6 | 16 | 18.44 | 16 | [5 empty] | 112
| 1 | 6 | 18.632 | 16.38 | 14.9 | 16 | 16.45 | 17.1 | 16.8 | 6 | 14 | 16.8 | 17.3654 | 15.1 | 16.15 | 14 | 17.1 | 17 | 18.39 | 17
| Mean |  | 17.8885 | 16.6067 | 15.45 | 15.9833 | 16.6583 | 16.8333 | 16.5483 | 6.5 | 14.3333 | 16.5583 | 17.7071 | 15.225 | 16.9333 | 14.6667 | 16.825 | 16.6667 | 18.3083 | 16.1667
| Median |  | 17.8125 | 16.565 | 15.3 | 15.95 | 16.675 | 16.85 | 16.545 | 6.5 | 14 | 16.65 | 17.7867 | 15.125 | 17.175 | 14.5 | 16.875 | 17 | 18.36 | 16
| Std Dev. |  | 0.48538 | 0.22214 | 0.50498 | 0.19664 | 0.28358 | 0.29439 | 0.33713 | 0.54772 | 0.5164 | 0.30235 | 0.33417 | 0.50473 | 0.51056 | 1.8619 | 0.40712 | 0.5164 | 0.35796 | 0.40825
| Rel.Std.Dev. |  | 2.713% | 1.338% | 3.268% | 1.23% | 1.702% | 1.749% | 2.037% | 8.427% | 3.603% | 1.826% | 1.887% | 3.315% | 3.015% | 12.69% | 2.42% | 3.098% | 1.955% | 2.525%
| PDM3 |  | 7.804% | 0.07889% | -6.892% | -3.678% | 0.3903% | 1.445% | -0.2727% | -60.83% | -13.62% | -0.2124% | 6.71% | -8.248% | 2.048% | -11.61% | 1.395% | 0.4405% | 10.33% | -2.573%
| Z-Score (Absolute) |  | 1.74 | 0.03 | 1.5 | 0.8 | 0.1 | 0.33 | 0.05 | 13.4 | 2.99 | 0.03 | 1.49 | 1.8 | 0.47 | 2.55 | 0.32 | 0.11 | 2.29 | 0.55
| Analytical results for Yb in OREAS 501b (Certified Value 1.5 ppm) | [24 empty] | Certified
Yb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab |  |  | 1
| No | No | 01: 02; 02: 03; 03: 07; 04: 08; 05: 16; 06: 23; 07: 25 |  |  | ppm
|  |  | 01: AR*OES/MS; 02: AR*MS; 03: AR*OES/MS; 04: AR*MS; 05: AR*OES/MS; 06: AR*OES/MS; 07: AR*OES/MS |  |  | 2
3
| 1 | 1 | 01: 1.43; 02: 1.29; 03: 1.4; 04: 1.52; 05: 1.512; 06: 1.48; 07: 1.5 |  |  | 1
| 1 | 2 | 01: 1; 02: 1.32; 03: 1.4; 04: 1.57; 05: 1.5249; 06: 1.47; 07: 1.5 |  |  | #N/A
| 1 | 3 | 01: 1.45; 02: 1.36; 03: 1.4; 04: 1.53; 05: 1.5469; 06: 1.49; 07: 1.4 |  |  | 16
| 1 | 4 | 01: 1.53; 02: 1.36; 03: 1.5; 04: 1.52; 05: 1.5586; 06: 1.49; 07: 1.5 |  |  | 1.4956
| 1 | 5 | 01: 1.47; 02: 1.32; 03: 1.4; 04: 1.58; 05: 1.5539; 06: 1.47; 07: 1.5 |  |  | 113
| 1 | 6 | 01: 1.53; 02: 1.28; 03: 1.5; 04: 1.57; 05: 1.5139; 06: 1.45; 07: 1.5
| Mean |  | 01: 1.4017; 02: 1.3217; 03: 1.4333; 04: 1.5483; 05: 1.535; 06: 1.475; 07: 1.4833
| Median |  | 01: 1.46; 02: 1.32; 03: 1.4; 04: 1.55; 05: 1.5359; 06: 1.475; 07: 1.5
| Std Dev. |  | 01: 0.20104; 02: 0.033714; 03: 0.05164; 04: 0.027869; 05: 0.020665; 06: 0.015166; 07: 0.040825
| Rel.Std.Dev. |  | 01: 14.34%; 02: 2.551%; 03: 3.603%; 04: 1.8%; 05: 1.346%; 06: 1.028%; 07: 2.752%
| PDM3 |  | 01: -6.282%; 02: -11.63%; 03: -4.164%; 04: 3.525%; 05: 2.635%; 06: -1.378%; 07: -0.8212%
| Z-Score (Absolute) |  | 01: 0.82; 02: 1.72; 03: 0.47; 04: 0.82; 05: 0.67; 06: 0; 07: 0.09
| Analytical results for Zn in OREAS 501b (Certified Value 80 ppm) | [24 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | AR*MS | AR*OES/MS | AR*OES | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*OES | AR*OES | AR*OES | AR*OES | AR*MS | AR*OES | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES | 0
1
| 1 | 1 | 87.756 | 81 | 78 | 76 | 80 | 74 | 74 | 90 | 80 | 86 | 87 | 78 | 78 | 75.0935 | 74 | 78 | 61 | 87 | 79.73 | 85 | 66 | 72 | 84 | 1
| 1 | 2 | 86.874 | 81 | 74 | 78 | 79 | 78 | 77 | 87 | 80 | 83 | 79 | 81 | 76 | 77.6992 | 75 | 77 | 64 | 89 | 79.8 | 83 | 66 | 79 | 76 | #N/A
| 1 | 3 | 85.747 | 82 | 78 | 75 | 79 | 71 | 76 | 87 | 78 | 87 | 77 | 76 | 75 | 76.3146 | 74 | 79 | 67 | 86 | 80.08 | 83 | 68 | 79 | 86 | 16
| 1 | 4 | 86.234 | 84 | 79 | 77 | 81 | 74 | 77 | 85 | 79 | 85 | 77 | 78 | 72 | 81.9473 | 76 | 80 | 81 | 83 | 80.05 | 85 | 65 | 82 | 86 | 79.8572
| 1 | 5 | 83.9 | 84 | 72 | 77 | 79 | 79 | 77 | 101 | 79 | 86 | 77 | 83 | 71 | 81.7728 | 77 | 76 | 70 | 81 | 79.34 | 84 | 64 | 84 | 80 | 114
| 1 | 6 | 87.684 | 81 | 74 | 78 | 79 | 75 | 77 | 85 | 76 | 85 | 78 | 82 | 72 | 76.4647 | 75 | 78 | 65 | 82 | 79.32 | 85 | 66 | 83 | 83
| Mean |  | 86.3658 | 82.1667 | 75.8333 | 76.8333 | 79.5 | 75.1667 | 76.3333 | 89.1667 | 78.6667 | 85.3333 | 79.1667 | 79.6667 | 74 | 78.2153 | 75.1667 | 78 | 68 | 84.6667 | 79.72 | 84.1667 | 65.8333 | 79.8333 | 82.5
| Median |  | 86.554 | 81.5 | 76 | 77 | 79 | 74.5 | 77 | 87 | 79 | 85.5 | 77.5 | 79.5 | 73.5 | 77.082 | 75 | 78 | 66 | 84.5 | 79.765 | 84.5 | 66 | 80.5 | 83.5
| Std Dev. |  | 1.4428 | 1.472 | 2.8577 | 1.169 | 0.83666 | 2.9269 | 1.2111 | 6.08 | 1.5055 | 1.3663 | 3.92 | 2.7325 | 2.7568 | 2.9419 | 1.169 | 1.4142 | 7.0427 | 3.1411 | 0.33148 | 0.98319 | 1.3292 | 4.3551 | 3.8859
| Rel.Std.Dev. |  | 1.671% | 1.791% | 3.768% | 1.522% | 1.052% | 3.894% | 1.587% | 6.819% | 1.914% | 1.601% | 4.952% | 3.43% | 3.725% | 3.761% | 1.555% | 1.813% | 10.36% | 3.71% | 0.4158% | 1.168% | 2.019% | 5.455% | 4.71%
| PDM3 |  | 8.15% | 2.892% | -5.039% | -3.787% | -0.4473% | -5.874% | -4.413% | 11.66% | -1.491% | 6.857% | -0.8647% | -0.2386% | -7.335% | -2.056% | -5.874% | -2.326% | -14.85% | 6.023% | -0.1718% | 5.396% | -17.56% | -0.02989% | 3.309%
| Z-Score (Absolute) |  | 1.46 | 0.61 | 0.67 | 0.47 | 0.07 | 0.81 | 0.57 | 2.02 | 0.1 | 1.25 | 0 | 0.1 | 1.05 | 0.19 | 0.81 | 0.24 | 2.26 | 1.11 | 0.11 | 1.01 | 2.7 | 0.13 | 0.67
| Analytical results for Zr in OREAS 501b (Certified Value 10.9 ppm) | [24 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 07 | 08 | 09 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*MS | AR*OES/MS | AR*MS | AR*OES | AR*OES | AR*MS | AR*OES/MS | AR*MS | AR*MS | AR*OES | AR*MS | AR*OES/MS | AR*OES/MS | AR*OES/MS | AR*OES | AR*OES |  |  |  | 1
1
| 1 | 1 | 9.9277 | 13.6 | 7 | 10.4 | 11 | 11.1 | 11.6 | 12.1 | 4 | 9.5 | 12.8239 | 10.5 | 12 | 9 | 9.1 | 10.05 | 10.3 | 8.6 | 25 | 12 |  |  |  | 1
| 1 | 2 | 10.333 | 13.8 | 8 | 10.9 | 11 | 11.3 | 12.2 | 12 | 4 | 9.1 | 13.5124 | 10.7 | 12.1 | 9 | 9.7 | 10.14 | 10 | 9 | 23 | 12 |  |  |  | #N/A
| 1 | 3 | 9.9654 | 14 | 8 | 10.2 | 11.3 | 11.8 | 11.8 | 12.1 | 4 | 9.3 | 12.7125 | 10.7 | 11.4 | 10 | 9.4 | 10.24 | 10.5 | 9.3 | 21 | 12 |  |  |  | 16
| 1 | 4 | 9.7538 | 14.8 | 8 | 10.5 | 11.2 | 11.6 | 11.9 | 12 | 4 | 9.6 | 13.0752 | 10.9 | 12.3 | 12 | 9.7 | 9.98 | 10.5 | 8.9 | 27 | 12 |  |  |  | 10.8518
| 1 | 5 | 9.8596 | 14.1 | 8 | 10.3 | 11.1 | 11.6 | 12.5 | 12.5 | 4 | 9.3 | 13.7988 | 11.4 | 12.2 | 10 | 9.6 | 10.02 | 10.1 | 8.6 | 21 | 12 |  |  |  | 115
| 1 | 6 | 9.9673 | 14.2 | 8 | 10.5 | 10.9 | 11.7 | 12.3 | 11.3 | 4 | 9.6 | 12.703 | 10.8 | 10.6 | 9 | 9.6 | 10.07 | 10.7 | 9.2 | 22 | 12
| Mean |  | 9.9678 | 14.0833 | 7.8333 | 10.4667 | 11.0833 | 11.5167 | 12.05 | 12 | 4 | 9.4 | 13.1043 | 10.8333 | 11.7667 | 9.8333 | 9.5167 | 10.0833 | 10.35 | 8.9333 | 23.1667 | 12
| Median |  | 9.9466 | 14.05 | 8 | 10.45 | 11.05 | 11.6 | 12.05 | 12.05 | 4 | 9.4 | 12.9496 | 10.75 | 12.05 | 9.5 | 9.6 | 10.06 | 10.4 | 8.95 | 22.5 | 12
| Std Dev. |  | 0.19621 | 0.41191 | 0.40825 | 0.24221 | 0.1472 | 0.26394 | 0.33912 | 0.38987 | 0 | 0.2 | 0.4567 | 0.30768 | 0.6532 | 1.169 | 0.23166 | 0.093524 | 0.26646 | 0.29439 | 2.4014 | 0
| Rel.Std.Dev. |  | 1.968% | 2.925% | 5.212% | 2.314% | 1.328% | 2.292% | 2.814% | 3.249% | 0% | 2.128% | 3.485% | 2.84% | 5.551% | 11.89% | 2.434% | 0.9275% | 2.574% | 3.295% | 10.37% | 0%
| PDM3 |  | -8.146% | 29.78% | -27.82% | -3.549% | 2.134% | 6.127% | 11.04% | 10.58% | -63.14% | -13.38% | 20.76% | -0.17% | 8.431% | -9.385% | -12.3% | -7.081% | -4.624% | -17.68% | 113.5% | 10.58%
| Z-Score (Absolute) |  | 0.39 | 1.94 | 1.59 | 0.1 | 0.25 | 0.49 | 0.79 | 0.76 | 3.76 | 0.71 | 1.39 | 0.1 | 0.63 | 0.46 | 0.64 | 0.32 | 0.17 | 0.97 | 7.08 | 0.76
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1151: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1169: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1187: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## 4-Acid
| Analytical results for Ag in OREAS 501b (Certified Value 0.778 ppm) | [26 empty] | Certified
Ag | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.7692 | 0.76 | 0.5 | 0.79 | 0.67 | 0.8 | < 2 | 0.69 | < 2 | 0.8 | 0.7 | 0.6 | < 0.5 | 0.9 | 0.77179 | 0.73 | 0.83 | 2 | 0.74 | 3.4 | 1 | 0.9 | 0.24 | 1.8 | 1.33 | 1
| 1 | 2 | 0.7548 | 0.78 | 0.5 | 0.74 | 0.66 | 0.8 | < 2 | 0.67 | < 2 | 0.9 | 0.8 | 0.64 | < 0.5 | 0.9 | 0.77243 | 0.71 | 0.87 | < 1 | 0.74 | 3.3 | 1.1 | 1.3 | 0.28 | 1.8 | 1.35 | #N/A
| 1 | 3 | 0.7643 | 0.75 | 1 | 0.76 | 0.68 | 0.7 | < 2 | 0.67 | < 2 | 0.8 | 0.6 | 0.64 | < 0.5 | 1.1 | 0.79623 | 0.7 | 0.82 | < 1 | 0.77 | 3.3 | 1 | 1.1 | 0.27 | 1.7 | 1.19 | 16
| 1 | 4 | 0.7669 | 0.82 | 1 | 0.78 | 0.7 | 0.8 | < 2 | 0.65 | 3 | 0.9 | 0.6 | 0.59 | < 0.5 | 0.8 | 0.7593 | 0.7 | 0.83 | < 1 | 0.77 | 3.5 | 0.9 | 0.8 | 0.4 | 1.9 | 1.42 | 0.77838
| 1 | 5 | 0.6741 | 0.78 | 0.5 | 0.72 | 0.7 | 0.9 | < 2 | 0.71 | < 2 | 0.9 | 0.7 | 0.66 | < 0.5 | 1.1 | 0.70974 | 0.73 | 0.83 | < 1 | 0.72 | 3.5 | 1 | 0.9 | 0.26 | 1.1 | 1.34 | 9
| 1 | 6 | 0.6848 | 0.76 | 0.5 | 0.82 | 0.6 | 0.8 | < 2 | 0.74 | < 2 | 0.9 | 0.8 | 0.64 | < 0.5 | 1 | 0.73883 | 0.68 | 0.83 | < 1 | 0.76 | 3.5 | 0.9 | 0.8 | 0.29 | 1.5 | 1.34
| Mean |  | 0.73568 | 0.775 | 0.66667 | 0.76833 | 0.66833 | 0.8 |  | 0.68833 | 3 | 0.86667 | 0.7 | 0.62833 |  | 0.96667 | 0.75805 | 0.70833 | 0.835 | 2 | 0.75 | 3.4167 | 0.98333 | 0.96667 | 0.29 | 1.6333 | 1.3283
| Median |  | 0.75955 | 0.77 | 0.5 | 0.77 | 0.675 | 0.8 |  | 0.68 | 3 | 0.9 | 0.7 | 0.64 |  | 0.95 | 0.76554 | 0.705 | 0.83 | 2 | 0.75 | 3.45 | 1 | 0.9 | 0.275 | 1.75 | 1.34
| Std Dev. |  | 0.043963 | 0.0251 | 0.2582 | 0.036009 | 0.037103 | 0.063246 |  | 0.032506 |  | 0.05164 | 0.089443 | 0.027142 |  | 0.12111 | 0.030194 | 0.019408 | 0.017607 |  | 0.02 | 0.098319 | 0.075277 | 0.19664 | 0.056569 | 0.29439 | 0.075211
| Rel.Std.Dev. |  | 5.976% | 3.239% | 38.73% | 4.687% | 5.552% | 7.906% |  | 4.722% |  | 5.958% | 12.78% | 4.32% |  | 12.53% | 3.983% | 2.74% | 2.109% |  | 2.667% | 2.878% | 7.655% | 20.34% | 19.51% | 18.02% | 5.662%
| PDM3 |  | -5.485% | -0.4338% | -14.35% | -1.29% | -14.14% | 2.778% |  | -11.57% | 285.4% | 11.34% | -10.07% | -19.28% |  | 24.19% | -2.611% | -8.999% | 7.275% | 156.9% | -3.646% | 338.9% | 26.33% | 24.19% | -62.74% | 109.8% | 70.65%
| Z-Score (Absolute) |  | 0.24 | 0 | 0.67 | 0.04 | 0.66 | 0.16 | 7.62 | 0.54 | 8.66 | 0.57 | 0.47 | 0.91 | 1.71 | 1.19 | 0.11 | 0.41 | 0.37 | 2.44 | 0.16 | 16.44 | 1.3 | 1.19 | 3.02 | 5.34 | 3.44
| Analytical results for Al in OREAS 501b (Certified Value 7.68 wt.%) | [26 empty] | Certified
Al | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 8.2114 | 8.057 | 7.98 | 7.45 | 7.54 | 7.55 | 8.18 | 7.5823 | 7.19 | 7.67 | 8.51 | 7.98 | 8.0789 | 7.96 | 8.094 | 7.53 | 7.5 | 3.61 | 7.58 | 7.8624 | 7.34 | 6.99 | 7.51 | 5.19 | 6.32 | 1
| 1 | 2 | 8.3378 | 8.029 | 7.85 | 7.33 | 7.54 | 7.32 | 8.24 | 7.5458 | 7.32 | 7.61 | 8.28 | 7.55 | 8.1801 | 7.97 | 8.143 | 7.2 | 8 | 3.06 | 7.49 | 7.5196 | 7.14 | 6.78 | 7.57 | 4.69 | 6.11 | #N/A
| 1 | 3 | 8.363 | 8.023 | 7.97 | 7.31 | 7.59 | 7.25 | 8.21 | 7.8135 | 7.35 | 7.39 | 8.33 | 7.32 | 8.0717 | 7.78 | 8.284 | 7.15 | 7.47 | 2.41 | 7.11 | 7.8561 | 7.4 | 6.83 | 7.78 | 4.79 | 5.85 | 16
| 1 | 4 | 8.3199 | 8.037 | 7.89 | 7.61 | 7.55 | 7.23 | 8.11 | 7.7655 | 7.54 | 7.67 | 6.81 | 7.31 | 8.0471 | 7.76 | 8.22 | 7.02 | 7.33 | 2.54 | 7.05 | 7.6979 | 7.07 | 6.7 | 7.5 | 4.87 | 6.6 | 7.678
| 1 | 5 | 8.6749 | 7.952 | 7.95 | 7.38 | 7.54 | 7.45 | 8.18 | 7.8418 | 7.52 | 7.57 | 8.37 | 7.62 | 8.003 | 7.9 | 8.19 | 7.4 | 7.38 | 4.53 | 7.7 | 7.6902 | 7.24 | 6.77 | 6.99 | 5.59 | 6.41 | 10
| 1 | 6 | 8.8646 | 7.927 | 7.88 | 7.77 | 7.64 | 7.14 | 8.16 | 7.8572 | 7.41 | 7.58 | 8.28 | 8.09 | 8.0235 | 8.05 | 7.934 | 7.06 | 7.45 | 3.41 | 7.24 | 7.702 | 7.17 | 7 | 6.8 | 4.71 | 6.2
| Mean |  | 8.4619 | 8.0042 | 7.92 | 7.475 | 7.5667 | 7.3233 | 8.18 | 7.7344 | 7.3883 | 7.5817 | 8.0967 | 7.645 | 8.0674 | 7.9033 | 8.1442 | 7.2267 | 7.5217 | 3.26 | 7.3617 | 7.7214 | 7.2267 | 6.845 | 7.3583 | 4.9733 | 6.2483
| Median |  | 8.3504 | 8.026 | 7.92 | 7.415 | 7.545 | 7.285 | 8.18 | 7.7895 | 7.38 | 7.595 | 8.305 | 7.585 | 8.0594 | 7.93 | 8.1665 | 7.175 | 7.46 | 3.235 | 7.365 | 7.7 | 7.205 | 6.805 | 7.505 | 4.83 | 6.26
| Std Dev. |  | 0.25125 | 0.051994 | 0.053666 | 0.18064 | 0.040825 | 0.15175 | 0.044272 | 0.13605 | 0.13136 | 0.10323 | 0.63604 | 0.32795 | 0.062191 | 0.11396 | 0.12173 | 0.19957 | 0.24244 | 0.77959 | 0.26604 | 0.127 | 0.12517 | 0.12341 | 0.37765 | 0.35246 | 0.25888
| Rel.Std.Dev. |  | 2.969% | 0.6496% | 0.6776% | 2.417% | 0.5395% | 2.072% | 0.5412% | 1.759% | 1.778% | 1.362% | 7.856% | 4.29% | 0.7709% | 1.442% | 1.495% | 2.762% | 3.223% | 23.91% | 3.614% | 1.645% | 1.732% | 1.803% | 5.132% | 7.087% | 4.143%
| PDM3 |  | 10.21% | 4.249% | 3.152% | -2.643% | -1.449% | -4.619% | 6.539% | 0.7345% | -3.772% | -1.254% | 5.453% | -0.4292% | 5.072% | 2.935% | 6.072% | -5.878% | -2.036% | -57.54% | -4.119% | 0.5654% | -5.878% | -10.85% | -4.163% | -35.23% | -18.62%
| Z-Score (Absolute) |  | 1.78 | 0.87 | 0.7 | 0.18 | 0 | 0.48 | 1.22 | 0.33 | 0.35 | 0.03 | 1.05 | 0.16 | 0.99 | 0.67 | 1.15 | 0.67 | 0.09 | 8.54 | 0.41 | 0.31 | 0.67 | 1.43 | 0.41 | 5.14 | 2.61
| Analytical results for As in OREAS 501b (Certified Value 17.9 ppm) | [26 empty] | Certified
As | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
1
| 1 | 1 | 21.814 | 22 | 17 | 19 | 12.9 | 18 | 22 | 15.3 | 50 | 22 | 17 | 20.7 | 31 | 17 | 16.7963 | 14.3 | 15.9 | 15 | 18.7 | 17 | 21 | 22 | 11 | 17 |  | 1
| 1 | 2 | 22.291 | 19.9 | 18 | 18.2 | 16.1 | 14 | 19 | 15.3 | < 35 | 17 | 13 | 17.3 | 29 | 12 | 20.7743 | 16.6 | 21.6 | 16 | 24 | 18 | 20 | 19 | 9 | 16 |  | #N/A
| 1 | 3 | 17.17 | 21.9 | 19 | 22.7 | 19.7 | 18 | 18 | 17.5 | < 35 | 20 | 17 | 16.6 | 36 | 11 | 18.5176 | 13.7 | 17.9 | 12 | 21 | 18 | 18 | 20 | 11 | 14 |  | 16
| 1 | 4 | 19.32 | 20.4 | 17 | 19.1 | 16.6 | 26 | 14 | 19.4 | < 35 | 18 | 10 | 20.4 | 36 | 14 | 15.7667 | 21.5 | 26.7 | 18 | 20.6 | 19 | 27 | 15 | 11 | 15 |  | 17.8609
| 1 | 5 | 18.85 | 20.1 | 19 | 16.6 | 13.9 | 16 | 20 | 20.4 | 50 | 18 | 15 | 13.1 | 34 | 12 | 20.1602 | 17.6 | 18.2 | 25 | 18.2 | 21 | 20 | 20 | 11 | 16 |  | 11
| 1 | 6 | 19.56 | 20.2 | 20 | 32 | 21.6 | 17 | 15 | 20.1 | < 35 | 19 | 12 | 15.4 | 31 | 22 | 17.7888 | 15.4 | 17.6 | 36 | 18.7 | 18 | 19 | 18 | 36 | 16
| Mean |  | 19.8342 | 20.75 | 18.3333 | 21.2667 | 16.8 | 18.1667 | 18 | 18 | 50 | 19 | 14 | 17.25 | 32.8333 | 14.6667 | 18.3006 | 16.5167 | 19.65 | 20.3333 | 20.2 | 18.5 | 20.8333 | 19 | 14.8333 | 15.6667
| Median |  | 19.44 | 20.3 | 18.5 | 19.05 | 16.35 | 17.5 | 18.5 | 18.45 | 50 | 18.5 | 14 | 16.95 | 32.5 | 13 | 18.1532 | 16 | 18.05 | 17 | 19.65 | 18 | 20 | 19.5 | 11 | 16
| Std Dev. |  | 1.9163 | 0.94393 | 1.2111 | 5.6262 | 3.3347 | 4.1191 | 3.0332 | 2.3221 | 0 | 1.7889 | 2.8284 | 2.929 | 2.9269 | 4.1793 | 1.9272 | 2.8323 | 3.9226 | 8.8242 | 2.179 | 1.3784 | 3.1885 | 2.3664 | 10.4003 | 1.0328
| Rel.Std.Dev. |  | 9.662% | 4.549% | 6.606% | 26.46% | 19.85% | 22.67% | 16.85% | 12.9% | 0% | 9.415% | 20.2% | 16.98% | 8.914% | 28.5% | 10.53% | 17.15% | 19.96% | 43.4% | 10.79% | 7.451% | 15.3% | 12.45% | 70.11% | 6.592%
| PDM3 |  | 11.05% | 16.18% | 2.645% | 19.07% | -5.94% | 1.712% | 0.779% | 0.779% | 179.9% | 6.378% | -21.62% | -3.42% | 83.83% | -17.88% | 2.462% | -7.526% | 10.02% | 13.84% | 13.1% | 3.578% | 16.64% | 6.378% | -16.95% | -12.28%
| Z-Score (Absolute) |  | 0.56 | 0.93 | 0.03 | 1.13 | 0.64 | 0.1 | 0.17 | 0.17 | 8.56 | 0.23 | 1.75 | 0.46 | 5.72 | 1.49 | 0.05 | 0.75 | 0.49 | 0.76 | 0.71 | 0.03 | 0.96 | 0.23 | 1.42 | 1.09
| Analytical results for B in OREAS 501b (Indicative Value 20 ppm) | [26 empty] | Indicative
B | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 03 |  | ppm
|  |  | 01: 4A*MS |  | 1
1
| 1 | 1 | 01: < 20 |  | 1
| 1 | 2 | 01: 20 |  | 3
| 1 | 3 | 01: 20 |  | 16
| 1 | 4 | 01: 20 |  | 20
| 1 | 5 | 01: 20 |  | 9
| 1 | 6 | 01: 20
| Mean |  | 01: 20
| Median |  | 01: 20
| Std Dev. |  | 01: 0
Rel.Std.Dev. | bar: 01
PDM3 | bar: 01
Z-Score (Absolute) | bar: 01
| Analytical results for Ba in OREAS 501b (Certified Value 1009 ppm) | [26 empty] | Certified
Ba | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 0
0
| 1 | 1 | 1036.29 | 1042.7 | 942 | 1010 | 950 | 1007 | 998 | 1068 | 989 | 1080 | 1153 | 1040 | 969 | 990 | 993.0881 | 1070 | 1090 | 480 | 1000 | 1043 | 879 | 1032 | 643 | 860 |  | 1
| 1 | 2 | 1014.74 | 1068.4 | 886 | 990 | 940 | 1026 | 975 | 1063 | 988 | 1050 | 1133 | 990 | 989 | 990 | 991.6005 | 1030 | 1140 | 439 | 1000 | 1044 | 873 | 1028 | 478 | 840 |  | #N/A
| 1 | 3 | 1055.63 | 1065 | 927 | 990 | 950 | 1007 | 979 | 1089.2 | 999 | 1020 | 1141 | 1020 | 961 | 970 | 1017.4328 | 1010 | 1080 | 401 | 1000 | 1039 | 866 | 1037 | 559 | 780 |  | 16
| 1 | 4 | 1036.47 | 1070.6 | 945 | 1030 | 950 | 1010 | 999 | 1084.5 | 1010 | 1020 | 965 | 1000 | 958 | 970 | 993.5342 | 1020 | 1070 | 415 | 1000 | 1040 | 858 | 990 | 450 | 820 |  | 1009.4709
| 1 | 5 | 947.892 | 1060.9 | 932 | 1000 | 950 | 1032 | 980 | 1083.6 | 1010 | 1050 | 1162 | 1010 | 950 | 980 | 986.882 | 1070 | 1080 | 530 | 1010 | 1065 | 894 | 981 | 432 | 880 |  | 12
| 1 | 6 | 999.508 | 1040.1 | 909 | 1050 | 960 | 1007 | 985 | 1076.3 | 1010 | 1060 | 1139 | 990 | 956 | 1000 | 987.5628 | 1010 | 1060 | 434 | 980 | 1028 | 862 | 927 | 421 | 840
| Mean |  | 1015.0883 | 1057.95 | 923.5 | 1011.6667 | 950 | 1014.8333 | 986 | 1077.4333 | 1001 | 1046.6667 | 1115.5 | 1008.3333 | 963.8333 | 983.3333 | 995.0167 | 1035 | 1086.6667 | 449.8333 | 998.3333 | 1043.1667 | 872 | 999.1667 | 497.1667 | 836.6667
| Median |  | 1025.515 | 1062.95 | 929.5 | 1005 | 950 | 1008.5 | 982.5 | 1079.95 | 1004.5 | 1050 | 1140 | 1005 | 959.5 | 985 | 992.3443 | 1025 | 1080 | 436.5 | 1000 | 1041.5 | 869.5 | 1009 | 464 | 840
| Std Dev. |  | 38.2323 | 13.2576 | 22.3853 | 24.0139 | 6.3246 | 11.1967 | 10.198 | 10.2471 | 10.583 | 23.3809 | 74.4681 | 19.4079 | 13.8191 | 12.1106 | 11.3296 | 28.1069 | 28.0476 | 47.5454 | 9.8319 | 12.123 | 13.1605 | 42.3387 | 86.8687 | 34.448
| Rel.Std.Dev. |  | 3.766% | 1.253% | 2.424% | 2.374% | 0.6657% | 1.103% | 1.034% | 0.9511% | 1.057% | 2.234% | 6.676% | 1.925% | 1.434% | 1.232% | 1.139% | 2.716% | 2.581% | 10.57% | 0.9848% | 1.162% | 1.509% | 4.237% | 17.47% | 4.117%
| PDM3 |  | 0.5565% | 4.802% | -8.516% | 0.2175% | -5.891% | 0.5312% | -2.325% | 6.732% | -0.8391% | 3.685% | 10.5% | -0.1127% | -4.521% | -2.589% | -1.432% | 2.529% | 7.647% | -55.44% | -1.103% | 3.338% | -13.62% | -1.021% | -50.75% | -17.12%
| Z-Score (Absolute) |  | 0.26 | 0.98 | 1.3 | 0.2 | 0.85 | 0.25 | 0.24 | 1.31 | 0.02 | 0.79 | 1.96 | 0.14 | 0.62 | 0.28 | 0.09 | 0.59 | 1.47 | 9.35 | 0.03 | 0.73 | 2.18 | 0.02 | 8.55 | 2.78
| Analytical results for Be in OREAS 501b (Certified Value 2.86 ppm) | [26 empty] | Certified
Be | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  | 2
2
| 1 | 1 | 3.6328 | 2.8 | 2.7 | 2.51 | 2.69 | 3 | 2.6 | 3 | 2.9 | 3 | 2.9 | 2.8 | 2.7995 | 3.21 | 3.08 | 2 | 2.81 | < 2 | 2.7 | 4 | 3.2 | 2.5 | < 5 |  |  | 1
| 1 | 2 | 3.5707 | 2.9 | 2.5 | 2.47 | 2.59 | 3 | 2.6 | 3.12 | 2.8 | 3 | 2.8 | 2.9 | 2.8182 | 3.17 | 3.39 | 3 | 2.95 | < 2 | 2.9 | 3.7 | 3.1 | 2.5 | < 5 |  |  | #N/A
| 1 | 3 | 3.4819 | 2.98 | 2.6 | 2.42 | 3.06 | 3 | 2.5 | 3.19 | 2.8 | 3 | 2.9 | 2.8 | 2.855 | 2.97 | 3.04 | 2 | 2.79 | < 2 | 2.9 | 4 | 2.6 | 2.3 | < 5 |  |  | 16
| 1 | 4 | 3.5894 | 3.22 | 2.8 | 2.74 | 2.53 | 3 | 2.6 | 3.23 | 2.9 | 3 | 2.8 | 2.8 | 2.8202 | 3.05 | 3.25 | 2 | 2.94 | < 2 | 2.8 | 4.1 | 3.2 | 2.7 | < 5 |  |  | 2.8567
| 1 | 5 | 2.9241 | 3.04 | 2.8 | 2.57 | 2.67 | 3 | 2.5 | 2.98 | 2.8 | 3 | 2.9 | 2.8 | 2.7516 | 3.18 | 3.06 | 3 | 2.9 | < 2 | 2.9 | 4.1 | 7.6 | 2.6 | < 5 |  |  | 13
| 1 | 6 | 2.9031 | 2.86 | 2.6 | 2.54 | 2.55 | 3 | 2.5 | 3.48 | 2.9 | 3 | 2.8 | 2.9 | 2.8246 | 3.11 | 3.08 | 2 | 2.98 | < 2 | 2.8 | 4.3 | 3.3 | 2.9 | < 5
| Mean |  | 3.3503 | 2.9667 | 2.6667 | 2.5417 | 2.6817 | 3 | 2.55 | 3.1667 | 2.85 | 3 | 2.85 | 2.8333 | 2.8115 | 3.115 | 3.15 | 2.3333 | 2.895 |  | 2.8333 | 4.0333 | 3.8333 | 2.5833
| Median |  | 3.5263 | 2.94 | 2.65 | 2.525 | 2.63 | 3 | 2.55 | 3.155 | 2.85 | 3 | 2.85 | 2.8 | 2.8192 | 3.14 | 3.08 | 2 | 2.92 |  | 2.85 | 4.05 | 3.2 | 2.55
| Std Dev. |  | 0.34191 | 0.15055 | 0.12111 | 0.11053 | 0.196 | 0 | 0.054772 | 0.18305 | 0.054772 | 0 | 0.054772 | 0.05164 | 0.034408 | 0.091159 | 0.13971 | 0.5164 | 0.078166 |  | 0.08165 | 0.19664 | 1.8619 | 0.20412
| Rel.Std.Dev. |  | 10.21% | 5.075% | 4.541% | 4.349% | 7.309% | 0% | 2.148% | 5.78% | 1.922% | 0% | 1.922% | 1.823% | 1.224% | 2.926% | 4.435% | 22.13% | 2.7% |  | 2.882% | 4.875% | 48.57% | 7.902%
| PDM3 |  | 17.28% | 3.848% | -6.654% | -11.03% | -6.129% | 5.015% | -10.74% | 10.85% | -0.236% | 5.015% | -0.236% | -0.8194% | -1.583% | 9.04% | 10.27% | -18.32% | 1.339% |  | -0.8194% | 41.19% | 34.19% | -9.571%
| Z-Score (Absolute) |  | 1.27 | 0.3 | 0.47 | 0.78 | 0.43 | 0.38 | 0.76 | 0.81 | 0 | 0.38 | 0 | 0.04 | 0.1 | 0.67 | 0.76 | 1.31 | 0.11 | 2.16 | 0.04 | 3.01 | 2.5 | 0.68 | 5.47
| Analytical results for Bi in OREAS 501b (Certified Value 1.54 ppm) | [26 empty] | Certified
Bi | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 1.7917 | 1.61 | 1.4 | 1.75 | 1.3 | 2.3 | 5 | 1.54 | < 30 | < 5 | < 5 | 1.6 | < 2 | < 2 | 1.6249 | 1.58 | 1.53 | < 1 | 1.42 | 1.6 | 1.32 | 1.53 | < 1 | < 5 |  | 1
| 1 | 2 | 1.6048 | 1.59 | 1.4 | 1.62 | 1.35 | 1.9 | 6 | 1.5 | < 30 | < 5 | < 5 | 1.59 | < 2 | < 2 | 1.6638 | 1.5 | 1.62 | < 1 | 1.5 | 2 | 1.35 | 1.44 | < 1 | < 5 |  | #N/A
| 1 | 3 | 1.5179 | 1.62 | 1.3 | 1.61 | 1.41 | 1.9 | < 5 | 1.61 | < 30 | < 5 | < 5 | 1.66 | < 2 | 2 | 1.644 | 1.56 | 1.59 | 2 | 1.62 | 1.7 | 1.37 | 1.46 | < 1 | < 5 |  | 16
| 1 | 4 | 1.6262 | 1.74 | 1.4 | 1.76 | 1.36 | 2.2 | < 5 | 1.65 | < 30 | < 5 | < 5 | 1.75 | < 2 | 2 | 1.6399 | 1.51 | 1.54 | < 1 | 1.49 | 1.7 | 1.27 | 1.49 | < 1 | < 5 |  | 1.5439
| 1 | 5 | 1.4819 | 1.54 | 1.4 | 1.6 | 1.29 | 2.1 | < 5 | 1.59 | < 30 | < 5 | < 5 | 1.64 | < 2 | 5 | 1.647 | 1.64 | 1.54 | < 1 | 1.35 | 1.8 | 1.34 | 1.54 | < 1 | < 5 |  | 14
| 1 | 6 | 1.3893 | 1.67 | 1.4 | 1.7 | 1.34 | 1.9 | < 5 | 1.59 | < 30 | < 5 | < 5 | 1.61 | < 2 | < 2 | 1.7089 | 1.49 | 1.52 | < 1 | 1.53 | 1.7 | 1.35 | 1.62 | < 1 | < 5
| Mean |  | 1.5686 | 1.6283 | 1.3833 | 1.6733 | 1.3417 | 2.05 | 5.5 | 1.58 |  |  |  | 1.6417 |  | 3 | 1.6548 | 1.5467 | 1.5567 | 2 | 1.485 | 1.75 | 1.3333 | 1.5133
| Median |  | 1.5614 | 1.615 | 1.4 | 1.66 | 1.345 | 2 | 5.5 | 1.59 |  |  |  | 1.625 |  | 2 | 1.6455 | 1.535 | 1.54 | 2 | 1.495 | 1.7 | 1.345 | 1.51
| Std Dev. |  | 0.13902 | 0.069113 | 0.040825 | 0.072572 | 0.043551 | 0.17607 | 0.70711 | 0.052915 |  |  |  | 0.059133 |  | 1.7321 | 0.029325 | 0.05785 | 0.039328 |  | 0.092682 | 0.13784 | 0.035024 | 0.065013
| Rel.Std.Dev. |  | 8.863% | 4.244% | 2.951% | 4.337% | 3.246% | 8.589% | 12.86% | 3.349% |  |  |  | 3.602% |  | 57.74% | 1.772% | 3.74% | 2.526% |  | 6.241% | 7.877% | 2.627% | 4.296%
| PDM3 |  | 1.603% | 5.47% | -10.4% | 8.385% | -13.1% | 32.78% | 256.2% | 2.339% |  |  |  | 6.333% |  | 94.31% | 7.181% | 0.1802% | 0.8279% | 29.54% | -3.814% | 13.35% | -13.64% | -1.979%
| Z-Score (Absolute) |  | 0.16 | 0.02 | 0.62 | 0.09 | 0.73 | 1.03 | 8.74 | 0.14 | 70.19 | 8.33 | 8.33 | 0.02 | 0.9 | 2.14 | 0.05 | 0.22 | 0.19 | 1.16 | 0.37 | 0.28 | 0.75 | 0.3 | 1.57 | 8.33
| Analytical results for Ca in OREAS 501b (Certified Value 2.71 wt.%) | [26 empty] | Certified
Ca | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 2.7677 | 2.802 | 2.85 | 2.7 | 2.8 | 2.77 | 2.52 | 2.6803 | 2.5 | 2.75 | 3.06 | 2.81 | 2.6489 | 2.69 | 2.731 | 2.83 | 2.79 | 1.14 | 2.75 | 2.2183 | 2.83 | 2.69 | 2.72 | 1.95 | 2.35 | 1
| 1 | 2 | 2.6859 | 2.875 | 2.87 | 2.64 | 2.79 | 2.74 | 2.52 | 2.6681 | 2.54 | 2.7 | 3.01 | 2.67 | 2.7023 | 2.7 | 2.751 | 2.68 | 2.9 | 0.97 | 2.7 | 2.1245 | 2.84 | 2.59 | 2.7 | 1.38 | 2.34 | #N/A
| 1 | 3 | 2.8388 | 2.859 | 2.85 | 2.67 | 2.79 | 2.67 | 2.44 | 2.7486 | 2.54 | 2.67 | 3.02 | 2.74 | 2.6122 | 2.63 | 2.754 | 2.69 | 2.76 | 0.86 | 2.68 | 2.2015 | 2.93 | 2.61 | 2.73 | 1.92 | 2.18 | 16
| 1 | 4 | 2.7303 | 2.851 | 2.87 | 2.83 | 2.78 | 2.74 | 2.52 | 2.7217 | 2.58 | 2.69 | 2.6 | 2.7 | 2.6879 | 2.63 | 2.715 | 2.73 | 2.74 | 0.9 | 2.71 | 2.1771 | 2.81 | 2.59 | 2.63 | 1.53 | 2.49 | 2.7108
| 1 | 5 | 2.5088 | 2.843 | 2.85 | 2.7 | 2.78 | 2.76 | 2.46 | 2.7514 | 2.56 | 2.71 | 3.08 | 2.71 | 2.6477 | 2.66 | 2.708 | 2.81 | 2.71 | 1.39 | 2.65 | 2.1557 | 2.88 | 2.62 | 2.54 | 1.67 | 2.37 | 15
| 1 | 6 | 2.6977 | 2.82 | 2.83 | 2.77 | 2.81 | 2.69 | 2.55 | 2.7783 | 2.57 | 2.74 | 2.98 | 2.65 | 2.7014 | 2.73 | 2.716 | 2.66 | 2.69 | 1.16 | 2.66 | 2.1961 | 2.79 | 2.7 | 2.52 | 1.51 | 2.39
| Mean |  | 2.7049 | 2.8417 | 2.8533 | 2.7183 | 2.7917 | 2.7283 | 2.5017 | 2.7247 | 2.5483 | 2.71 | 2.9583 | 2.7133 | 2.6667 | 2.6733 | 2.7292 | 2.7333 | 2.765 | 1.07 | 2.6917 | 2.1789 | 2.8467 | 2.6333 | 2.64 | 1.66 | 2.3533
| Median |  | 2.714 | 2.847 | 2.85 | 2.7 | 2.79 | 2.74 | 2.52 | 2.7351 | 2.55 | 2.705 | 3.015 | 2.705 | 2.6684 | 2.675 | 2.7235 | 2.71 | 2.75 | 1.055 | 2.69 | 2.1866 | 2.835 | 2.615 | 2.665 | 1.6 | 2.36
| Std Dev. |  | 0.1108 | 0.026621 | 0.015055 | 0.06969 | 0.01169 | 0.039707 | 0.042151 | 0.043223 | 0.028577 | 0.030332 | 0.17916 | 0.056804 | 0.036246 | 0.040332 | 0.01959 | 0.071181 | 0.075033 | 0.1992 | 0.03656 | 0.034234 | 0.050859 | 0.049261 | 0.092304 | 0.23221 | 0.10053
| Rel.Std.Dev. |  | 4.096% | 0.9368% | 0.5276% | 2.564% | 0.4188% | 1.455% | 1.685% | 1.586% | 1.121% | 1.119% | 6.056% | 2.094% | 1.359% | 1.509% | 0.7178% | 2.604% | 2.714% | 18.62% | 1.358% | 1.571% | 1.787% | 1.871% | 3.496% | 13.99% | 4.272%
| PDM3 |  | -0.218% | 4.829% | 5.259% | 0.2788% | 2.984% | 0.6477% | -7.714% | 0.5149% | -5.992% | -0.02859% | 9.132% | 0.09438% | -1.625% | -1.381% | 0.6785% | 0.8322% | 2% | -60.53% | -0.7049% | -19.62% | 5.013% | -2.857% | -2.611% | -38.76% | -13.19%
| Z-Score (Absolute) |  | 0.05 | 1.27 | 1.38 | 0.08 | 0.79 | 0.18 | 2.01 | 0.14 | 1.56 | 0 | 2.39 | 0.03 | 0.42 | 0.35 | 0.18 | 0.22 | 0.53 | 15.8 | 0.18 | 5.12 | 1.32 | 0.74 | 0.67 | 10.11 | 3.44
| Analytical results for Cd in OREAS 501b (Indicative Value 0.15 ppm) | [26 empty] | Indicative
Cd | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 2
2
| 1 | 1 | 0.2645 | 0.13 | < 0.5 | 0.15 | 0.14 | < 0.1 | 1 | 0.26 | < 1 | < 1 | < 0.4 | 0.06 | 12.4 | < 0.5 | 0.17728 | 0.09 | 0.16 | 6 | 0.1 | < 2 | < 0.2 | 0.22 | 0.03 | 5 | 0.62 | 1
| 1 | 2 | 0.2222 | 0.11 | < 0.5 | 0.15 | 0.15 | 0.1 | 1 | 0.24 | < 1 | < 1 | < 0.4 | 0.08 | 12.8 | < 0.5 | 0.16147 | 0.09 | 0.18 | 4 | 0.11 | < 2 | < 0.2 | 0.22 | 0.08 | 5 | < 0.5 | 4
| 1 | 3 | 0.2259 | 0.12 | < 0.5 | 0.21 | 0.14 | < 0.1 | 1 | 0.22 | < 1 | < 1 | < 0.4 | 0.07 | 11.9 | < 0.5 | 0.15964 | 0.11 | 0.14 | 4 | 0.11 | < 2 | < 0.2 | 0.22 | 0.03 | 5 | < 0.5 | 16
| 1 | 4 | 0.2403 | 0.15 | < 0.5 | 0.15 | 0.12 | < 0.1 | 1 | 0.2 | < 1 | < 1 | < 0.4 | 0.08 | 12.7 | < 0.5 | 0.15949 | 0.09 | 0.15 | 5 | 0.16 | < 2 | < 0.2 | 0.22 | 0.17 | 5 | 0.53 | 0.14556
| 1 | 5 | 0.2419 | 0.12 | < 0.5 | 0.14 | 0.13 | < 0.1 | 1 | 0.26 | < 1 | < 1 | < 0.4 | 0.05 | 8.8 | < 0.5 | 0.13483 | 0.1 | 0.14 | 4 | 0.14 | < 2 | < 0.2 | 0.23 | 0.57 | 4 | 0.52 | 10
| 1 | 6 | 0.2339 | 0.1 | < 0.5 | 0.2 | 0.13 | < 0.1 | 1 | 0.18 | < 1 | < 1 | < 0.4 | 0.1 | 9.9 | < 0.5 | 0.13032 | 0.11 | 0.15 | 5 | 0.11 | < 2 | < 0.2 | 0.23 | 0.13 | 5 | < 0.5
| Mean |  | 0.23812 | 0.12167 |  | 0.16667 | 0.135 | 0.1 | 1 | 0.22667 |  |  |  | 0.073333 | 11.4167 |  | 0.15384 | 0.098333 | 0.15333 | 4.6667 | 0.12167 |  |  | 0.22333 | 0.16833 | 4.8333 | 0.55667
| Median |  | 0.2371 | 0.12 |  | 0.15 | 0.135 | 0.1 | 1 | 0.23 |  |  |  | 0.075 | 12.15 |  | 0.15957 | 0.095 | 0.15 | 4.5 | 0.11 |  |  | 0.22 | 0.105 | 5 | 0.53
| Std Dev. |  | 0.01507 | 0.017224 |  | 0.030111 | 0.010488 |  | 0 | 0.03266 |  |  |  | 0.017512 | 1.6678 |  | 0.017819 | 0.0098319 | 0.015055 | 0.8165 | 0.023166 |  |  | 0.005164 | 0.20439 | 0.40825 | 0.055076
| Rel.Std.Dev. |  | 6.329% | 14.16% |  | 18.07% | 7.769% |  | 0% | 14.41% |  |  |  | 23.88% | 14.61% |  | 11.58% | 9.999% | 9.819% | 17.5% | 19.04% |  |  | 2.312% | 121.4% | 8.447% | 9.894%
| PDM3 |  | 63.59% | -16.42% |  | 14.5% | -7.255% | -31.3% | 587% | 55.72% |  |  |  | -49.62% | 7743% |  | 5.687% | -32.45% | 5.34% | 3106% | -16.42% |  |  | 53.43% | 15.64% | 3220% | 282.4%
| Z-Score (Absolute) |  | 0.06 | 0.55 | 1.44 | 0.32 | 0.48 | 0.67 | 4.06 | 0 | 4.06 | 4.06 | 0.91 | 0.81 | 58.8 | 1.44 | 0.38 | 0.67 | 0.39 | 23.33 | 0.55 | 9.32 | 0.14 | 0.02 | 0.31 | 24.21 | 1.59
| Analytical results for Ce in OREAS 501b (Certified Value 68 ppm) | [26 empty] | Certified
Ce | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [8 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [8 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [8 empty] | 0
1
| 1 | 1 | 74.093 | 69.55 | 62.1 | 74.6 | 59.3 | 63 | 72.32 | 60.7 | 80 | 69.375 | 71 | 69.7 | 50 | 64.2 | 72.4 | 64.4 | 68.99 | [8 empty] | 1
| 1 | 2 | 72.348 | 73.43 | 57.6 | 70.9 | 58.1 | 63 | 68.2 | 60.2 | 76 | 66.028 | 66.5 | 67.7 | 51 | 63.2 | 70.1 | 65.3 | 76.86 | [8 empty] | #N/A
| 1 | 3 | 77.359 | 72.92 | 60.2 | 77.1 | 66.4 | 65 | 73.17 | 60.6 | 76 | 65.497 | 66.8 | 60.9 | 46 | 61.5 | 73 | 67.8 | 63.06 | [8 empty] | 16
| 1 | 4 | 76.44 | 73.89 | 60.2 | 78.9 | 62.3 | 63 | 69.11 | 60.7 | 81 | 68.581 | 62.2 | 62.9 | 46 | 66.5 | 72 | 65.5 | 62.11 | [8 empty] | 67.629
| 1 | 5 | 66.672 | 72.68 | 60.4 | 74.4 | 59.8 | 63 | 68.58 | 61.2 | 76 | 65.301 | 66.1 | 61.2 | 53 | 66.2 | 74.1 | 68.3 | 55.59 | [8 empty] | 16
| 1 | 6 | 67.997 | 72.43 | 57.4 | 82.4 | 66.3 | 63 | 72.17 | 58.7 | 78 | 65.193 | 67.4 | 71.7 | 51 | 57.5 | 71.3 | 69 | 65.04
| Mean |  | 72.4848 | 72.4833 | 59.65 | 76.3833 | 62.0333 | 63.3333 | 70.5917 | 60.35 | 77.8333 | 66.6625 | 66.6667 | 65.6833 | 49.5 | 63.1833 | 72.15 | 66.7167 | 65.275
| Median |  | 73.2205 | 72.8 | 60.2 | 75.85 | 61.05 | 63 | 70.64 | 60.65 | 77 | 65.7625 | 66.65 | 65.3 | 50.5 | 63.7 | 72.2 | 66.65 | 64.05
| Std Dev. |  | 4.3803 | 1.5305 | 1.813 | 4.0057 | 3.6131 | 0.8165 | 2.1949 | 0.86891 | 2.2286 | 1.8337 | 2.8154 | 4.6288 | 2.881 | 3.3546 | 1.3809 | 1.8841 | 7.1665
| Rel.Std.Dev. |  | 6.043% | 2.111% | 3.039% | 5.244% | 5.824% | 1.289% | 3.109% | 1.44% | 2.863% | 2.751% | 4.223% | 7.047% | 5.82% | 5.309% | 1.914% | 2.824% | 10.98%
| PDM3 |  | 7.18% | 7.178% | -11.8% | 12.94% | -8.274% | -6.352% | 4.381% | -10.76% | 15.09% | -1.429% | -1.423% | -2.877% | -26.81% | -6.574% | 6.685% | -1.349% | -3.481%
| Z-Score (Absolute) |  | 0.85 | 0.85 | 1.02 | 1.42 | 0.67 | 0.48 | 0.57 | 0.92 | 1.63 | 0 | 0 | 0.14 | 2.5 | 0.51 | 0.8 | 0.01 | 0.2
| Analytical results for Co in OREAS 501b (Certified Value 15.8 ppm) | [26 empty] | Certified
Co | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 18.891 | 16.5 | 20 | 15.3 | 14.2 | 17.8 | 14 | 16.3 | 14 | 17 | 15 | 16 | 14 | 15 | 16.51 | 16.4 | 16 | 17 | 15.1 | 15 | 16 | 19.2 | 17.3 | 20 | 22 | 1
| 1 | 2 | 19.32 | 17.3 | 20 | 14.9 | 14 | 17.6 | 14 | 17 | 15 | 17 | 14 | 16 | 14 | 15 | 16.08 | 16.2 | 17 | 14 | 16.1 | 14 | 16 | 15.5 | 16.7 | 20 | 22 | #N/A
| 1 | 3 | 18.168 | 17.5 | 20 | 15.7 | 14.1 | 16.8 | 14 | 17.2 | 15 | 17 | 15 | 15 | 14 | 12 | 16.09 | 16 | 15.5 | 16 | 16.2 | 15 | 17 | 16 | 16.2 | 20 | 21 | 16
| 1 | 4 | 18.864 | 17.9 | 20 | 16.4 | 14 | 17.5 | 14 | 17.7 | 15 | 19 | 14 | 15 | 14 | 14 | 16.11 | 16.3 | 16.8 | 16 | 16.7 | 15 | 16 | 16 | 15.3 | 21 | 24 | 15.8251
| 1 | 5 | 16.644 | 16.8 | 20 | 14.9 | 14 | 17.4 | 13 | 16.9 | 15 | 15 | 15 | 16 | 15 | 13 | 16.1 | 17.1 | 16 | 17 | 16.1 | 15 | 17 | 16.4 | 16 | 20 | 23 | 17
| 1 | 6 | 17.082 | 17.5 | 20 | 16.1 | 14 | 17.6 | 13 | 18.1 | 15 | 15 | 14 | 16 | 14 | 14 | 16.25 | 15.8 | 15.6 | 18 | 16.2 | 15 | 16 | 16.8 | 17.2 | 20 | 22
| Mean |  | 18.1615 | 17.25 | 20 | 15.55 | 14.05 | 17.45 | 13.6667 | 17.2 | 14.8333 | 16.6667 | 14.5 | 15.6667 | 14.1667 | 13.8333 | 16.19 | 16.3 | 16.15 | 16.3333 | 16.0667 | 14.8333 | 16.3333 | 16.65 | 16.45 | 20.1667 | 22.3333
| Median |  | 18.516 | 17.4 | 20 | 15.5 | 14 | 17.55 | 14 | 17.1 | 15 | 17 | 14.5 | 16 | 14 | 14 | 16.105 | 16.25 | 16 | 16.5 | 16.15 | 15 | 16 | 16.2 | 16.45 | 20 | 22
| Std Dev. |  | 1.0804 | 0.51284 | 0 | 0.6253 | 0.083666 | 0.34496 | 0.5164 | 0.63246 | 0.40825 | 1.5055 | 0.54772 | 0.5164 | 0.40825 | 1.169 | 0.16888 | 0.44721 | 0.61887 | 1.3663 | 0.52409 | 0.40825 | 0.5164 | 1.3233 | 0.76616 | 0.40825 | 1.0328
| Rel.Std.Dev. |  | 5.949% | 2.973% | 0% | 4.021% | 0.5955% | 1.977% | 3.779% | 3.677% | 2.752% | 9.033% | 3.777% | 3.296% | 2.882% | 8.451% | 1.043% | 2.744% | 3.832% | 8.365% | 3.262% | 2.752% | 3.162% | 7.947% | 4.658% | 2.024% | 4.624%
| PDM3 |  | 14.76% | 9.004% | 26.38% | -1.738% | -11.22% | 10.27% | -13.64% | 8.688% | -6.267% | 5.318% | -8.373% | -1.001% | -10.48% | -12.59% | 2.306% | 3.001% | 2.053% | 3.212% | 1.527% | -6.267% | 3.212% | 5.213% | 3.949% | 27.43% | 41.13%
| Z-Score (Absolute) |  | 1.32 | 0.67 | 2.63 | 0.53 | 1.6 | 0.82 | 1.87 | 0.64 | 1.04 | 0.26 | 1.28 | 0.45 | 1.51 | 1.75 | 0.08 | 0 | 0.11 | 0.02 | 0.17 | 1.04 | 0.02 | 0.25 | 0.11 | 2.74 | 4.28
| Analytical results for Cr in OREAS 501b (Certified Value 86 ppm) | [26 empty] | Certified
Cr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 94.8174 | 107 | 100 | 84 | 87 | 91 | 75 | 80 | 108 | 95 | 92 | 90 | 82 | 83 | 78.9094 | 88 | 90 | < 1 | 93 | 49 | 91 | 70 | 84 | 74 | 163 | 1
| 1 | 2 | 95.5457 | 105 | 100 | 89 | 89 | 91 | 78 | 74 | 112 | 90 | 95 | 84 | 79 | 84 | 76.9921 | 80 | 95 | < 1 | 89 | 38 | 89 | 68 | 91 | 74 | 164 | #N/A
| 1 | 3 | 97.1161 | 129 | 100 | 85 | 82 | 89 | 72 | 94 | 119 | 87 | 89 | 89 | 77 | 82 | 78.5277 | 80 | 88 | < 1 | 87 | 41 | 93 | 70 | 81 | 76 | 152 | 16
| 1 | 4 | 101.365 | 99 | 150 | 90 | 89 | 90 | 74 | 83 | 105 | 96 | 79 | 88 | 77 | 83 | 76.7178 | 82 | 112 | < 1 | 85 | 45 | 88 | 71 | 82 | 93 | 159 | 85.5038
| 1 | 5 | 94.25 | 100 | 150 | 85 | 85 | 95 | 78 | 79 | 109 | 87 | 92 | 88 | 86 | 87 | 74.3986 | 85 | 86 | < 1 | 88 | 42 | 86 | 67 | 69 | 91 | 157 | 18
| 1 | 6 | 111.36 | 99 | 100 | 89 | 88 | 90 | 81 | 83 | 108 | 94 | 87 | 86 | 75 | 84 | 77.1968 | 79 | 90 | < 1 | 86 | 42 | 96 | 71 | 67 | 82 | 159
| Mean |  | 99.0757 | 106.5 | 116.6667 | 87 | 86.6667 | 91 | 76.3333 | 82.1667 | 110.1667 | 91.5 | 89 | 87.5 | 79.3333 | 83.8333 | 77.1237 | 82.3333 | 93.5 |  | 88 | 42.8333 | 90.5 | 69.5 | 79 | 81.6667 | 159
| Median |  | 96.3309 | 102.5 | 100 | 87 | 87.5 | 90.5 | 76.5 | 81.5 | 108.5 | 92 | 90.5 | 88 | 78 | 83.5 | 77.0944 | 81 | 90 |  | 87.5 | 42 | 90 | 70 | 81.5 | 79 | 159
| Std Dev. |  | 6.5401 | 11.5195 | 25.8199 | 2.6077 | 2.7325 | 2.0976 | 3.266 | 6.6758 | 4.8751 | 4.0373 | 5.6214 | 2.1679 | 4.0332 | 1.7224 | 1.5983 | 3.5024 | 9.5446 |  | 2.8284 | 3.7639 | 3.6194 | 1.6432 | 9.2304 | 8.5479 | 4.3359
| Rel.Std.Dev. |  | 6.601% | 10.82% | 22.13% | 2.997% | 3.153% | 2.305% | 4.279% | 8.125% | 4.425% | 4.412% | 6.316% | 2.478% | 5.084% | 2.055% | 2.072% | 4.254% | 10.21% |  | 3.214% | 8.787% | 3.999% | 2.364% | 11.68% | 10.47% | 2.727%
| PDM3 |  | 15.87% | 24.56% | 36.45% | 1.75% | 1.36% | 6.428% | -10.73% | -3.903% | 28.84% | 7.013% | 4.089% | 2.335% | -7.217% | -1.954% | -9.801% | -3.708% | 9.352% |  | 2.919% | -49.9% | 5.843% | -18.72% | -7.606% | -4.488% | 85.96%
| Z-Score (Absolute) |  | 1.25 | 2.02 | 3.08 | 0 | 0.03 | 0.41 | 1.11 | 0.5 | 2.4 | 0.47 | 0.21 | 0.05 | 0.8 | 0.33 | 1.02 | 0.48 | 0.67 | 8.92 | 0.1 | 4.58 | 0.36 | 1.82 | 0.83 | 0.55 | 7.47
| Analytical results for Cs in OREAS 501b (Certified Value 12.2 ppm) | [26 empty] | Certified
Cs | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab; 10: Lab; 11: Lab; 12: Lab; 13: Lab; 15: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 04; 05: 05; 06: 08; 07: 12; 08: 16; 09: 17; 10: 18; 11: 20; 12: 23; 13: 24; 15: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS; 09: 4A*MS; 10: 4A*MS; 11: 4A*MS; 12: 4A*MS; 13: 4A*MS; 15: 4A*OES/MS | 1
2
| 1 | 1 | 01: 12.566; 02: 11.44; 03: 9.6; 04: 12.6; 05: 10.1; 06: 11.9; 07: 10.8; 08: 13.0754; 09: 12.75; 10: 11.8; 11: 10.65; 12: 11.8; 13: 12.9; 15: 12 | 1
| 1 | 2 | 01: 12.514; 02: 11.82; 03: 9.2; 04: 11.85; 05: 9.76; 06: 12.26; 07: 11.1; 08: 13.0614; 09: 12.35; 10: 12.85; 11: 11.1; 12: 12.3; 13: 13.1; 15: 12 | #N/A
| 1 | 3 | 01: 12.061; 02: 11.86; 03: 9.5; 04: 12.3; 05: 10.35; 06: 12.35; 07: 11.5; 08: 13.2097; 09: 12.2; 10: 11.65; 11: 11.75; 12: 11.1; 13: 13; 15: 12 | 16
| 1 | 4 | 01: 12.54; 02: 12; 03: 9.9; 04: 12.8; 05: 10.4; 06: 12.46; 07: 11.2; 08: 13.2143; 09: 12.2; 10: 12.2; 11: 11.7; 12: 12; 13: 12.7; 15: 11 | 12.1535
| 1 | 5 | 01: 11.116; 02: 12.2; 03: 10; 04: 12.05; 05: 9.9; 06: 12.33; 07: 11.3; 08: 12.9532; 09: 12.6; 10: 11.6; 11: 11; 12: 12.1; 13: 13.2; 15: 14 | 19
| 1 | 6 | 01: 10.965; 02: 11.94; 03: 9.5; 04: 13.1; 05: 10.7; 06: 12.48; 07: 11.3; 08: 13.1708; 09: 12.35; 10: 11.9; 11: 10.8; 12: 12.1; 13: 13.5; 15: 13
| Mean |  | 01: 11.9603; 02: 11.8767; 03: 9.6167; 04: 12.45; 05: 10.2017; 06: 12.2967; 07: 11.2; 08: 13.1141; 09: 12.4083; 10: 12; 11: 11.1667; 12: 11.9; 13: 13.0667; 15: 12.3333
| Median |  | 01: 12.2875; 02: 11.9; 03: 9.55; 04: 12.45; 05: 10.225; 06: 12.34; 07: 11.25; 08: 13.1231; 09: 12.35; 10: 11.85; 11: 11.05; 12: 12.05; 13: 13.05; 15: 12
| Std Dev. |  | 01: 0.73799; 02: 0.25216; 03: 0.29269; 04: 0.47117; 05: 0.34816; 06: 0.21116; 07: 0.23664; 08: 0.10253; 09: 0.2223; 10: 0.46797; 11: 0.46007; 12: 0.42426; 13: 0.27325; 15: 1.0328
| Rel.Std.Dev. |  | 01: 6.17%; 02: 2.123%; 03: 3.044%; 04: 3.784%; 05: 3.413%; 06: 1.717%; 07: 2.113%; 08: 0.7818%; 09: 1.792%; 10: 3.9%; 11: 4.12%; 12: 3.565%; 13: 2.091%; 15: 8.374%
| PDM3 |  | 01: -1.589%; 02: -2.278%; 03: -20.87%; 04: 2.44%; 05: -16.06%; 06: 1.178%; 07: -7.846%; 08: 7.904%; 09: 2.097%; 10: -1.263%; 11: -8.12%; 12: -2.086%; 13: 7.514%; 15: 1.48%
| Z-Score (Absolute) |  | 01: 0.03; 02: 0.16; 03: 3.55; 04: 0.71; 05: 2.67; 06: 0.48; 07: 1.17; 08: 1.7; 09: 0.64; 10: 0.03; 11: 1.22; 12: 0.12; 13: 1.63; 15: 0.53
| Analytical results for Cu in OREAS 501b (Certified Value 0.26 wt.%) | [26 empty] | Certified
Cu | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.27402 | 0.2696 | 0.264 | 0.263 | 0.258 | 0.2638 | 0.262 | 0.2588 | 0.234 | 0.268 | 0.2945 | 0.274 | 0.2579 | 0.272 | 0.26554 | 0.262 | 0.268 | 0.2179 | 0.268 | 0.2419 | 0.2582 | 0.2371 | 0.31324 | 0.23 | 0.2174 | 1
| 1 | 2 | 0.26504 | 0.2709 | 0.263 | 0.257 | 0.258 | 0.26208 | 0.264 | 0.25797 | 0.238 | 0.263 | 0.2837 | 0.259 | 0.2635 | 0.277 | 0.26678 | 0.25 | 0.279 | 0.2056 | 0.266 | 0.2348 | 0.2572 | 0.2285 | 0.30179 | 0.24 | 0.2141 | #N/A
| 1 | 3 | 0.27242 | 0.2725 | 0.262 | 0.259 | 0.259 | 0.25802 | 0.262 | 0.26805 | 0.236 | 0.258 | 0.2872 | 0.254 | 0.2597 | 0.266 | 0.2772 | 0.248 | 0.264 | 0.2076 | 0.259 | 0.2336 | 0.2638 | 0.2284 | 0.3064 | 0.24 | 0.2016 | 16
| 1 | 4 | 0.27299 | 0.2804 | 0.27 | 0.277 | 0.257 | 0.2582 | 0.262 | 0.2631 | 0.244 | 0.265 | 0.2531 | 0.259 | 0.2542 | 0.269 | 0.27651 | 0.255 | 0.265 | 0.208 | 0.265 | 0.243 | 0.2495 | 0.2286 | 0.28464 | 0.25 | 0.2279 | 0.25967
| 1 | 5 | 0.25151 | 0.2617 | 0.26 | 0.262 | 0.256 | 0.26087 | 0.263 | 0.26832 | 0.239 | 0.261 | 0.289 | 0.265 | 0.2526 | 0.269 | 0.27084 | 0.265 | 0.264 | 0.2264 | 0.256 | 0.2378 | 0.253 | 0.2291 | 0.28241 | 0.25 | 0.2188 | 20
| 1 | 6 | 0.26084 | 0.2727 | 0.255 | 0.271 | 0.26 | 0.25634 | 0.262 | 0.26906 | 0.244 | 0.26 | 0.2824 | 0.273 | 0.259 | 0.278 | 0.26331 | 0.246 | 0.257 | 0.2214 | 0.262 | 0.2429 | 0.2537 | 0.2369 | 0.27352 | 0.25 | 0.2188
| Mean |  | 0.26614 | 0.2713 | 0.26233 | 0.26483 | 0.258 | 0.25988 | 0.2625 | 0.26422 | 0.23917 | 0.2625 | 0.28165 | 0.264 | 0.25782 | 0.27183 | 0.27003 | 0.25433 | 0.26617 | 0.21448 | 0.26267 | 0.239 | 0.2559 | 0.23143 | 0.29367 | 0.24333 | 0.21643
| Median |  | 0.26873 | 0.2717 | 0.2625 | 0.2625 | 0.258 | 0.25954 | 0.262 | 0.26558 | 0.2385 | 0.262 | 0.28545 | 0.262 | 0.25845 | 0.2705 | 0.26881 | 0.2525 | 0.2645 | 0.21295 | 0.2635 | 0.23985 | 0.25545 | 0.22885 | 0.29322 | 0.245 | 0.2181
| Std Dev. |  | 0.008851 | 0.0060236 | 0.0049261 | 0.0076529 | 0.0014142 | 0.0028284 | 0.00083666 | 0.0049919 | 0.0041191 | 0.0036194 | 0.014627 | 0.0081486 | 0.0039382 | 0.0047924 | 0.0058301 | 0.0077632 | 0.0072503 | 0.0086006 | 0.0045461 | 0.0041919 | 0.0049719 | 0.0043191 | 0.015655 | 0.008165 | 0.0085918
| Rel.Std.Dev. |  | 3.326% | 2.22% | 1.878% | 2.89% | 0.5481% | 1.088% | 0.3187% | 1.889% | 1.722% | 1.379% | 5.193% | 3.087% | 1.528% | 1.763% | 2.159% | 3.052% | 2.724% | 4.01% | 1.731% | 1.754% | 1.943% | 1.866% | 5.331% | 3.355% | 3.97%
| PDM3 |  | 2.491% | 4.479% | 1.026% | 1.989% | -0.6429% | 0.08304% | 1.09% | 1.751% | -7.896% | 1.09% | 8.465% | 1.668% | -0.7135% | 4.684% | 3.99% | -2.055% | 2.502% | -17.4% | 1.154% | -7.96% | -1.452% | -10.87% | 13.09% | -6.291% | -16.65%
| Z-Score (Absolute) |  | 0.37 | 0.9 | 0.02 | 0.24 | 0.46 | 0.27 | 0 | 0.18 | 2.38 | 0 | 1.96 | 0.15 | 0.48 | 0.95 | 0.77 | 0.83 | 0.37 | 4.91 | 0.02 | 2.4 | 0.67 | 3.17 | 3.18 | 1.96 | 4.71
| Analytical results for Dy in OREAS 501b (Certified Value 4.86 ppm) | [26 empty] | Certified
Dy | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | 2
3
| 1 | 1 | 01: 4.8028; 02: 4.99; 03: 4.3; 04: 4.73; 05: 4.8; 06: 4.81 | 1
| 1 | 2 | 01: 4.6983; 02: 4.99; 03: 4; 04: 4.78; 05: 4.9; 06: 4.73 | 9
| 1 | 3 | 01: 4.924; 02: 5.13; 03: 4.15; 04: 4.93; 05: 4.7; 06: 4.76 | 16
| 1 | 4 | 01: 4.9261; 02: 5.51; 03: 4.15; 04: 4.77; 05: 4.9; 06: 4.72 | 4.8596
| 1 | 5 | 01: 4.2689; 02: 4.89; 03: 4.1; 04: 4.73; 05: 5.1; 06: 4.9 | 21
| 1 | 6 | 01: 4.6212; 02: 5.22; 03: 4; 04: 4.77; 05: 5; 06: 4.96
| Mean |  | 01: 4.7069; 02: 5.1217; 03: 4.1167; 04: 4.785; 05: 4.9; 06: 4.8133
| Median |  | 01: 4.7505; 02: 5.06; 03: 4.125; 04: 4.77; 05: 4.9; 06: 4.785
| Std Dev. |  | 01: 0.24643; 02: 0.22311; 03: 0.11255; 04: 0.074229; 05: 0.14142; 06: 0.097502
| Rel.Std.Dev. |  | 01: 5.236%; 02: 4.356%; 03: 2.734%; 04: 1.551%; 05: 2.886%; 06: 2.026%
| PDM3 |  | 01: -3.142%; 02: 5.393%; 03: -15.29%; 04: -1.535%; 05: 0.8318%; 06: -0.9516%
| Z-Score (Absolute) |  | 01: 0.64; 02: 2.25; 03: 4.77; 04: 0.1; 05: 0.7; 06: 0.1
| Analytical results for Er in OREAS 501b (Certified Value 2.58 ppm) | [26 empty] | Certified
Er | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | 2
3
| 1 | 1 | 01: 2.467; 02: 2.79; 03: 2.4; 04: 2.68; 05: 2.8; 06: 2.49 | 1
| 1 | 2 | 01: 2.4151; 02: 2.72; 03: 2.3; 04: 2.73; 05: 2.6; 06: 2.5 | 10
| 1 | 3 | 01: 2.5024; 02: 2.77; 03: 2.4; 04: 2.83; 05: 2.7; 06: 2.51 | 16
| 1 | 4 | 01: 2.4661; 02: 2.86; 03: 2.3; 04: 2.68; 05: 2.9; 06: 2.48 | 2.5816
| 1 | 5 | 01: 2.2497; 02: 2.79; 03: 2.35; 04: 2.64; 05: 2.8; 06: 2.54 | 22
| 1 | 6 | 01: 2.3252; 02: 2.73; 03: 2.3; 04: 2.68; 05: 2.8; 06: 2.59
| Mean |  | 01: 2.4043; 02: 2.7767; 03: 2.3417; 04: 2.7067; 05: 2.7667; 06: 2.5183
| Median |  | 01: 2.4406; 02: 2.78; 03: 2.325; 04: 2.68; 05: 2.8; 06: 2.505
| Std Dev. |  | 01: 0.09762; 02: 0.050465; 03: 0.04916; 04: 0.066833; 05: 0.10328; 06: 0.040702
| Rel.Std.Dev. |  | 01: 4.06%; 02: 1.817%; 03: 2.099%; 04: 2.469%; 05: 3.733%; 06: 1.616%
| PDM3 |  | 01: -6.87%; 02: 7.556%; 03: -9.294%; 04: 4.845%; 05: 7.169%; 06: -2.451%
| Z-Score (Absolute) |  | 01: 0.88; 02: 0.7; 03: 1.15; 04: 0.4; 05: 0.65; 06: 0.4
| Analytical results for Eu in OREAS 501b (Certified Value 1.39 ppm) | [26 empty] | Certified
Eu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | 2
3
| 1 | 1 | 01: 1.4732; 02: 1.41; 03: 1.25; 04: 1.31; 05: 1.4; 06: 1.33 | 1
| 1 | 2 | 01: 1.4238; 02: 1.4; 03: 1.15; 04: 1.31; 05: 1.5; 06: 1.35 | 11
| 1 | 3 | 01: 1.4887; 02: 1.39; 03: 1.2; 04: 1.37; 05: 1.3; 06: 1.36 | 16
| 1 | 4 | 01: 1.4637; 02: 1.45; 03: 1.15; 04: 1.33; 05: 1.4; 06: 1.35 | 1.3872
| 1 | 5 | 01: 1.3229; 02: 1.36; 03: 1.2; 04: 1.29; 05: 1.5; 06: 1.38 | 23
| 1 | 6 | 01: 1.3844; 02: 1.37; 03: 1.15; 04: 1.29; 05: 1.5; 06: 1.41
| Mean |  | 01: 1.4261; 02: 1.3967; 03: 1.1833; 04: 1.3167; 05: 1.4333; 06: 1.3633
| Median |  | 01: 1.4438; 02: 1.395; 03: 1.175; 04: 1.31; 05: 1.45; 06: 1.355
| Std Dev. |  | 01: 0.063159; 02: 0.032042; 03: 0.040825; 04: 0.030111; 05: 0.08165; 06: 0.028048
| Rel.Std.Dev. |  | 01: 4.429%; 02: 2.294%; 03: 3.45%; 04: 2.287%; 05: 5.696%; 06: 2.057%
| PDM3 |  | 01: 2.804%; 02: 0.6807%; 03: -14.7%; 04: -5.086%; 05: 3.324%; 06: -1.722%
| Z-Score (Absolute) |  | 01: 0.63; 02: 0.23; 03: 2.67; 04: 0.86; 05: 0.72; 06: 0.23
| Analytical results for Fe in OREAS 501b (Certified Value 4.54 wt.%) | [26 empty] | Certified
Fe | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 4.615 | 4.58 | 4.79 | 4.5 | 4.48 | 4.5 | 4.26 | 4.62 | 4.1 | 4.56 | 4.93 | 4.66 | 4.65 | 4.63 | 4.621 | 4.58 | 4.61 | 3.72 | 4.68 | 5.3358 | 4.82 | 4.28 | 4.54 | 4.05 | 3.96 | 1
| 1 | 2 | 4.5819 | 4.66 | 4.75 | 4.42 | 4.46 | 4.48 | 4.18 | 4.69 | 4.15 | 4.56 | 4.81 | 4.43 | 4.76 | 4.62 | 4.625 | 4.39 | 4.84 | 3.54 | 4.67 | 4.8382 | 4.81 | 4.15 | 4.62 | 3.93 | 4.03 | #N/A
| 1 | 3 | 4.7533 | 4.67 | 4.74 | 4.43 | 4.48 | 4.4 | 4.11 | 4.72 | 4.17 | 4.43 | 4.86 | 4.54 | 4.64 | 4.55 | 4.76 | 4.35 | 4.56 | 3.55 | 4.59 | 5.0407 | 4.73 | 4.18 | 4.46 | 3.87 | 3.7 | 16
| 1 | 4 | 4.6704 | 4.65 | 4.68 | 4.52 | 4.47 | 4.47 | 4.26 | 4.7 | 4.2 | 4.65 | 4.29 | 4.46 | 4.61 | 4.54 | 4.718 | 4.34 | 4.55 | 3.57 | 4.56 | 5.1882 | 4.82 | 4.15 | 4.6 | 4.02 | 4.2 | 4.543
| 1 | 5 | 4.5498 | 4.61 | 4.76 | 4.38 | 4.47 | 4.5 | 4.11 | 4.64 | 4.18 | 4.47 | 4.94 | 4.49 | 4.58 | 4.57 | 4.644 | 4.55 | 4.53 | 3.92 | 4.69 | 4.9857 | 4.89 | 4.17 | 5.12 | 4.31 | 4.02 | 24
| 1 | 6 | 4.5506 | 4.58 | 4.7 | 4.62 | 4.55 | 4.42 | 4.34 | 4.75 | 4.2 | 4.54 | 4.87 | 4.44 | 4.67 | 4.71 | 4.606 | 4.31 | 4.47 | 3.8 | 4.53 | 4.9453 | 4.71 | 4.29 | 4.49 | 4.28 | 4.01
| Mean |  | 4.6202 | 4.625 | 4.7367 | 4.4783 | 4.485 | 4.4617 | 4.21 | 4.6867 | 4.1667 | 4.535 | 4.7833 | 4.5033 | 4.6517 | 4.6033 | 4.6623 | 4.42 | 4.5933 | 3.6833 | 4.62 | 5.0556 | 4.7967 | 4.2033 | 4.6383 | 4.0767 | 3.9867
| Median |  | 4.5984 | 4.63 | 4.745 | 4.465 | 4.475 | 4.475 | 4.22 | 4.695 | 4.175 | 4.55 | 4.865 | 4.475 | 4.645 | 4.595 | 4.6345 | 4.37 | 4.555 | 3.645 | 4.63 | 5.0132 | 4.815 | 4.175 | 4.57 | 4.035 | 4.015
| Std Dev. |  | 0.07941 | 0.040373 | 0.040332 | 0.086814 | 0.032711 | 0.042151 | 0.09252 | 0.048854 | 0.037771 | 0.077136 | 0.24639 | 0.08641 | 0.061779 | 0.06377 | 0.062047 | 0.11559 | 0.1291 | 0.15629 | 0.068702 | 0.17924 | 0.066232 | 0.064395 | 0.24384 | 0.18107 | 0.16244
| Rel.Std.Dev. |  | 1.719% | 0.8729% | 0.8515% | 1.939% | 0.7293% | 0.9447% | 2.198% | 1.042% | 0.9065% | 1.701% | 5.151% | 1.919% | 1.328% | 1.385% | 1.331% | 2.615% | 2.811% | 4.243% | 1.487% | 3.545% | 1.381% | 1.532% | 5.257% | 4.442% | 4.075%
| PDM3 |  | 1.699% | 1.806% | 4.264% | -1.423% | -1.276% | -1.79% | -7.329% | 3.163% | -8.283% | -0.1756% | 5.291% | -0.8726% | 2.392% | 1.329% | 2.627% | -2.707% | 1.108% | -18.92% | 1.695% | 11.28% | 5.584% | -7.476% | 2.099% | -10.26% | -12.25%
| Z-Score (Absolute) |  | 0.16 | 0.19 | 0.84 | 0.67 | 0.64 | 0.77 | 2.25 | 0.55 | 2.5 | 0.34 | 1.11 | 0.53 | 0.34 | 0.06 | 0.4 | 1.02 | 0 | 5.34 | 0.16 | 2.71 | 1.19 | 2.29 | 0.26 | 3.03 | 3.56
| Analytical results for Ga in OREAS 501b (Certified Value 19.5 ppm) | [26 empty] | Certified
Ga | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 08 | 10 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | [6 empty] | 1
2
| 1 | 1 | 20.949 | 19.58 | 17.6 | 17.4 | 16 | 19.06 | 20 | 20 | 20 | 19.4398 | 18.9 | 19.2 | 20 | 18.4 | 19 | 18.5 | 20.3 | 34 | 21 | [6 empty] | 1
| 1 | 2 | 20.508 | 19.94 | 16.2 | 16.85 | 16.2 | 19.24 | 20 | 20 | 20 | 19.0963 | 18.5 | 21 | 19 | 18.75 | 19.8 | 18.1 | 19.8 | 30 | 23 | [6 empty] | #N/A
| 1 | 3 | 20.362 | 20.04 | 16.6 | 17.4 | 16.95 | 19.38 | 20 | 20 | 10 | 19.0362 | 18.25 | 19.35 | 18 | 19.55 | 19.4 | 18.6 | 18.7 | 31 | 18 | [6 empty] | 16
| 1 | 4 | 20.334 | 20.31 | 17.4 | 18.4 | 16.7 | 19.82 | 20 | 20 | 20 | 19.275 | 18.25 | 20.2 | 17 | 20.5 | 19.8 | 18.4 | 18.1 | 28 | 24 | [6 empty] | 19.5236
| 1 | 5 | 17.77 | 19.42 | 17 | 17.35 | 16.25 | 19.57 | 20 | 20 | 20 | 19.2313 | 19.15 | 19.5 | 21 | 19 | 19.7 | 18.9 | 21.6 | 31 | 23 | [6 empty] | 25
| 1 | 6 | 18.239 | 20.55 | 16.4 | 18.55 | 16.75 | 20.21 | 20 | 20 | 20 | 19.2202 | 18.15 | 19.5 | 20 | 18.1 | 19.8 | 19.5 | 20.3 | 25 | 23
| Mean |  | 19.6937 | 19.9733 | 16.8667 | 17.6583 | 16.475 | 19.5467 | 20 | 20 | 18.3333 | 19.2164 | 18.5333 | 19.7917 | 19.1667 | 19.05 | 19.5833 | 18.6667 | 19.8 | 29.8333 | 22
| Median |  | 20.348 | 19.99 | 16.8 | 17.4 | 16.475 | 19.475 | 20 | 20 | 20 | 19.2257 | 18.375 | 19.5 | 19.5 | 18.875 | 19.75 | 18.55 | 20.05 | 30.5 | 23
| Std Dev. |  | 1.3351 | 0.42716 | 0.56095 | 0.6674 | 0.37517 | 0.41807 | 0 | 0 | 4.0825 | 0.14174 | 0.40579 | 0.68441 | 1.472 | 0.86833 | 0.32506 | 0.48442 | 1.2522 | 3.0605 | 2.1909
| Rel.Std.Dev. |  | 6.779% | 2.139% | 3.326% | 3.779% | 2.277% | 2.139% | 0% | 0% | 22.27% | 0.7376% | 2.19% | 3.458% | 7.68% | 4.558% | 1.66% | 2.595% | 6.324% | 10.26% | 9.959%
| PDM3 |  | 0.8712% | 2.304% | -13.61% | -9.554% | -15.61% | 0.1183% | 2.44% | 2.44% | -6.096% | -1.573% | -5.072% | 1.373% | -1.828% | -2.426% | 0.3061% | -4.389% | 1.416% | 52.81% | 12.68%
| Z-Score (Absolute) |  | 0.22 | 0.63 | 3.99 | 2.81 | 4.57 | 0 | 0.67 | 0.67 | 1.8 | 0.49 | 1.51 | 0.36 | 0.57 | 0.74 | 0.05 | 1.31 | 0.38 | 15.3 | 3.65
| Analytical results for Gd in OREAS 501b (Certified Value 5.4 ppm) | [26 empty] | Certified
Gd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 08: Lab | [6 empty] | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 08: 24 | [6 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 08: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 01: 5.1556; 02: 5.41; 03: 4.6; 04: 5.25; 05: 5.4; 08: 5.47 | [6 empty] | 1
| 1 | 2 | 01: 5.3425; 02: 5.4; 03: 4.4; 04: 5.19; 05: 5.5; 08: 5.52 | [6 empty] | 7
| 1 | 3 | 01: 5.6715; 02: 5.4; 03: 4.6; 04: 5.32; 05: 5.5; 08: 5.47 | [6 empty] | 16
| 1 | 4 | 01: 5.4004; 02: 5.3; 03: 4.6; 04: 5.25; 05: 5.6; 08: 5.38 | [6 empty] | 5.3977
| 1 | 5 | 01: 4.8401; 02: 5.38; 03: 4.4; 04: 5.27; 05: 5.4; 08: 5.51 | [6 empty] | 26
| 1 | 6 | 01: 5.2963; 02: 5.4; 03: 4.2; 04: 5.22; 05: 5.5; 08: 5.65
| Mean |  | 01: 5.2844; 02: 5.3817; 03: 4.4667; 04: 5.25; 05: 5.4833; 08: 5.5
| Median |  | 01: 5.3194; 02: 5.4; 03: 4.5; 04: 5.25; 05: 5.5; 08: 5.49
| Std Dev. |  | 01: 0.27598; 02: 0.041191; 03: 0.1633; 04: 0.044272; 05: 0.075277; 08: 0.088544
| Rel.Std.Dev. |  | 01: 5.222%; 02: 0.7654%; 03: 3.656%; 04: 0.8433%; 05: 1.373%; 08: 1.61%
| PDM3 |  | 01: -2.098%; 02: -0.2962%; 03: -17.25%; 04: -2.735%; 05: 1.587%; 08: 1.896%
| Z-Score (Absolute) |  | 01: 0.28; 02: 0.28; 03: 5.01; 04: 0.48; 05: 0.87; 08: 0.97
| Analytical results for Ge in OREAS 501b (Indicative Value 0.18 ppm) | [26 empty] | Indicative
Ge | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [12 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 08 | 12 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [12 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [12 empty] | 2
2
| 1 | 1 | 1.58 | < 20 | 0.16 | 0.17 | 1.04 | 0.24 | 0.18 | 0.24 | < 1 | 0.18 | 1.5 | < 0.1 | 0.1 | [12 empty] | 1
| 1 | 2 | 1.43 | < 20 | 0.16 | 0.16 | 1.15 | 0.24 | 0.18 | 0.32 | < 1 | 0.16 | 1.5 | < 0.1 | 0.1 | [12 empty] | 5
| 1 | 3 | 1.54 | < 20 | 0.16 | 0.18 | 1.26 | 0.26 | 0.17 | 0.22 | < 1 | 0.13 | 1.3 | < 0.1 | 0.2 | [12 empty] | 16
| 1 | 4 | 1.57 | < 20 | 0.18 | 0.17 | 1.28 | 0.26 | 0.16 | 0.27 | < 1 | 0.12 | 1.5 | < 0.1 | 0.2 | [12 empty] | 0.17958
| 1 | 5 | 1.46 | < 20 | 0.18 | 0.17 | 1.11 | 0.25 | 0.18 | 0.24 | 6 | 0.11 | 1.5 | 0.1 | 2.2 | [12 empty] | 11
| 1 | 6 | 1.48 | < 20 | 0.18 | 0.16 | 1.19 | 0.24 | 0.17 | 0.26 | < 1 | 0.13 | 1.5 | < 0.1 | 0.3
| Mean |  | 1.51 |  | 0.17 | 0.16833 | 1.1717 | 0.24833 | 0.17333 | 0.25833 | 6 | 0.13833 | 1.4667 | 0.1 | 0.51667
| Median |  | 1.51 |  | 0.17 | 0.17 | 1.17 | 0.245 | 0.175 | 0.25 | 6 | 0.13 | 1.5 | 0.1 | 0.2
| Std Dev. |  | 01: 0.061968; 03: 0.010954; 04: 0.0075277; 05: 0.091086; 08: 0.0098319; 10: 0.008165; 12: 0.034881; 16: 0.026394; 17: 0.08165; 19: 0.82805
| Rel.Std.Dev. |  | 01: 4.104%; 03: 6.444%; 04: 4.472%; 05: 7.774%; 08: 3.959%; 10: 4.711%; 12: 13.5%; 16: 19.08%; 17: 5.567%; 19: 160.3%
| PDM3 |  | 740.8% |  | -5.336% | -6.265% | 552.4% | 38.28% | -3.48% | 43.85% | 3241% | -22.97% | 716.7% | -44.32% | 187.7%
| Z-Score (Absolute) |  | 5.33 | 84.08 | 0.38 | 0.38 | 3.89 | 0.04 | 0.36 | 0 | 6.71 | 0.51 | 5.15 | 0.67 | 1.1
| Analytical results for Hf in OREAS 501b (Certified Value 2.54 ppm) | [26 empty] | Certified
Hf | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | 3.0375 | 2.89 | 2.4 | 2.6 | 2.4 | 2.1 | 2.84 | 2.8 | 2.6396 | 2.7 | 2.4 | 1 | 2.5 | 2.6 | 2.17 | 3.21 | [9 empty] | 1
| 1 | 2 | 2.99 | 2.77 | 2.4 | 2.5 | 2.2 | 2.3 | 2.56 | 2.8 | 2.4108 | 2.9 | 2.8 | < 1 | 2.2 | 2.7 | 2.02 | 2.75 | [9 empty] | #N/A
| 1 | 3 | 2.9287 | 2.71 | 2.2 | 2.6 | 2.3 | 2.1 | 2.56 | 2.7 | 2.473 | 2.8 | 2.6 | < 1 | 2.4 | 2.6 | 2.08 | 2.83 | [9 empty] | 16
| 1 | 4 | 2.9371 | 2.7 | 2.4 | 2.6 | 2.3 | 2.1 | 2.46 | 2.9 | 2.6458 | 2.8 | 2.4 | < 1 | 2.5 | 2.4 | 2.08 | 2.6 | [9 empty] | 2.5448
| 1 | 5 | 8.2496 | 2.66 | 2.2 | 2.5 | 2.2 | 2.3 | 2.51 | 2.8 | 2.3904 | 2.8 | 2.5 | 1 | 2.4 | 2.6 | 2.05 | 2.74 | [9 empty] | 27
| 1 | 6 | 3.3091 | 2.61 | 2.2 | 2.9 | 2.3 | 2.3 | 2.65 | 2.6 | 2.3826 | 2.8 | 2.5 | 1 | 2.2 | 2.5 | 2.26 | 3.17
| Mean |  | 3.9087 | 2.7233 | 2.3 | 2.6167 | 2.2833 | 2.2 | 2.5967 | 2.7667 | 2.4904 | 2.8 | 2.5333 | 1 | 2.3667 | 2.5667 | 2.11 | 2.8833
| Median |  | 3.0137 | 2.705 | 2.3 | 2.6 | 2.3 | 2.2 | 2.56 | 2.8 | 2.4419 | 2.8 | 2.5 | 1 | 2.4 | 2.6 | 2.08 | 2.79
| Std Dev. |  | 2.1312 | 0.097502 | 0.10954 | 0.1472 | 0.075277 | 0.10954 | 0.13486 | 0.10328 | 0.1222 | 0.063246 | 0.15055 | 0 | 0.13663 | 0.10328 | 0.088994 | 0.24913
| Rel.Std.Dev. |  | 54.53% | 3.58% | 4.763% | 5.625% | 3.297% | 4.979% | 5.194% | 3.733% | 4.907% | 2.259% | 5.943% | 0% | 5.773% | 4.024% | 4.218% | 8.64%
| PDM3 |  | 53.59% | 7.015% | -9.62% | 2.824% | -10.27% | -13.55% | 2.038% | 8.718% | -2.14% | 10.03% | -0.451% | -60.7% | -7% | 0.8588% | -17.09% | 13.3%
| Z-Score (Absolute) |  | 3.93 | 0.5 | 0.72 | 0.19 | 0.77 | 1.01 | 0.13 | 0.63 | 0.17 | 0.72 | 0.05 | 4.48 | 0.53 | 0.05 | 1.27 | 0.96
| Analytical results for Hg in OREAS 501b (Indicative Value 1.2 ppm) | [26 empty] | Indicative
Hg | Round | Replicate | 01: Lab | [9 empty] | 1
| No | No | 01: 19 | [9 empty] | ppm
|  |  | 01: 4A*OES | [9 empty] | 2
2
| 1 | 1 | 01: < 1 | [9 empty] | 1
| 1 | 2 | 01: 2 | [9 empty] | 3
| 1 | 3 | 01: < 1 | [9 empty] | 16
| 1 | 4 | 01: < 1 | [9 empty] | 1.2
| 1 | 5 | 01: 3 | [9 empty] | 9
| 1 | 6 | 01: < 1
| Mean |  | 01: 2.5
| Median |  | 01: 2.5
| Std Dev. |  | 01: 0.70711
| Rel.Std.Dev. |  | 01: 28.28%
| PDM3 |  | 01: 108.3%
Z-Score (Absolute) | bar: 01
| Analytical results for Ho in OREAS 501b (Certified Value 0.93 ppm) | [26 empty] | Certified
Ho | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [9 empty] | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | [9 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | [9 empty] | 2
3
| 1 | 1 | 01: 0.9558; 02: 0.89; 03: 0.86; 04: 0.93; 05: 0.9; 06: 0.84 | [9 empty] | 1
| 1 | 2 | 01: 0.9246; 02: 0.94; 03: 0.8; 04: 0.94; 05: 0.9; 06: 0.86 | [9 empty] | 9
| 1 | 3 | 01: 0.999; 02: 0.96; 03: 0.8; 04: 0.98; 05: 1; 06: 0.86 | [9 empty] | 16
| 1 | 4 | 01: 0.9448; 02: 0.94; 03: 0.82; 04: 0.94; 05: 1; 06: 0.85 | [9 empty] | 0.92701
| 1 | 5 | 01: 0.8464; 02: 0.97; 03: 0.82; 04: 0.93; 05: 1; 06: 0.88 | [9 empty] | 28
| 1 | 6 | 01: 0.9297; 02: 0.96; 03: 0.82; 04: 0.96; 05: 0.9; 06: 0.88
| Mean |  | 01: 0.93338; 02: 0.94333; 03: 0.82; 04: 0.94667; 05: 0.95; 06: 0.86167
| Median |  | 01: 0.93725; 02: 0.95; 03: 0.82; 04: 0.94; 05: 0.95; 06: 0.86
| Std Dev. |  | 01: 0.05019; 02: 0.028752; 03: 0.021909; 04: 0.019664; 05: 0.054772; 06: 0.016021
| Rel.Std.Dev. |  | 01: 5.377%; 02: 3.048%; 03: 2.672%; 04: 2.077%; 05: 5.766%; 06: 1.859%
| PDM3 |  | 01: 0.6875%; 02: 1.761%; 03: -11.54%; 04: 2.12%; 05: 2.48%; 06: -7.049%
| Z-Score (Absolute) |  | 01: 0.34; 02: 0.34; 03: 8; 04: 0.56; 05: 0.79; 06: 5.18
| Analytical results for In in OREAS 501b (Certified Value 0.21 ppm) | [26 empty] | Certified
In | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 2
2
| 1 | 1 | < 0.5 | 0.207 | 0.18 | 0.221 | 0.188 | 0.23 | 0.226 | 0.22 | 0.24388 | 0.237 | 0.2 | 14 | 0.256 | 0.2 | 0.52 | 0.22 | [9 empty] | 1
| 1 | 2 | < 0.5 | 0.229 | 0.16 | 0.224 | 0.185 | 0.22 | 0.214 | 0.216 | 0.24551 | 0.22 | 0.224 | 15 | 0.295 | 0.2 | 0.51 | 0.2 | [9 empty] | #N/A
| 1 | 3 | < 0.5 | 0.208 | 0.16 | 0.22 | 0.203 | 0.14 | 0.215 | 0.224 | 0.24393 | 0.216 | 0.198 | 20 | 0.239 | 0.2 | 0.5 | 0.2 | [9 empty] | 16
| 1 | 4 | < 0.5 | 0.21 | 0.18 | 0.235 | 0.19 | 0.15 | 0.227 | 0.215 | 0.24853 | 0.227 | 0.222 | 16 | 0.239 | 0.2 | 0.5 | 0.18 | [9 empty] | 0.21222
| 1 | 5 | < 0.5 | 0.206 | 0.18 | 0.221 | 0.196 | 0.24 | 0.223 | 0.213 | 0.24157 | 0.233 | 0.203 | 16 | 0.21 | 0.2 | 0.53 | 0.25 | [9 empty] | 29
| 1 | 6 | < 0.5 | 0.21 | 0.18 | 0.238 | 0.209 | 0.19 | 0.231 | 0.224 | 0.23752 | 0.219 | 0.199 | 21 | 0.22 | 0.2 | 0.56 | 0.21
| Mean |  |  | 0.21167 | 0.17333 | 0.2265 | 0.19517 | 0.195 | 0.22267 | 0.21867 | 0.24349 | 0.22533 | 0.20767 | 17 | 0.24317 | 0.2 | 0.52 | 0.21
| Median |  |  | 0.209 | 0.18 | 0.2225 | 0.193 | 0.205 | 0.2245 | 0.218 | 0.2439 | 0.2235 | 0.2015 | 16 | 0.239 | 0.2 | 0.515 | 0.205
| Std Dev. |  |  | 0.008641 | 0.010328 | 0.0079183 | 0.0093256 | 0.042308 | 0.0068313 | 0.0047188 | 0.0037198 | 0.0084063 | 0.012011 | 2.8284 | 0.030076 | 3.0405e-17 | 0.022804 | 0.023664
| Rel.Std.Dev. |  |  | 4.082% | 5.958% | 3.496% | 4.778% | 21.7% | 3.068% | 2.158% | 1.528% | 3.731% | 5.784% | 16.64% | 12.37% | 1.52e-14% | 4.385% | 11.27%
| PDM3 |  |  | -0.2594% | -18.32% | 6.73% | -8.034% | -8.113% | 4.924% | 3.039% | 14.74% | 6.181% | -2.144% | 7911% | 14.58% | -5.757% | 145% | -1.045%
| Z-Score (Absolute) |  | 8.73 | 0.28 | 1.48 | 0.18 | 0.8 | 0.8 | 0.06 | 0.06 | 0.71 | 0.15 | 0.41 | 524.22 | 0.7 | 0.65 | 9.35 | 0.33
| Analytical results for K in OREAS 501b (Certified Value 3.14 wt.%) | [26 empty] | Certified
K | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 3.1915 | 3.195 | 3.02 | 3.05 | 3.06 | 3.3 | 3.2 | 3.2893 | 3.29 | 2.94 | 3.07 | 3.19 | 3.2204 | 3.01 | 3.1141 | 3.25 | 3.24 | 0.22 | 3.15 | 3.8811 | 3.49 | 2.81 | 3.46 | 2.58 | 0.85 | 1
| 1 | 2 | 3.1636 | 3.304 | 3.03 | 2.98 | 3.07 | 3.24 | 2.84 | 3.3043 | 3.37 | 2.93 | 2.96 | 3.05 | 3.3073 | 3.07 | 3.1426 | 3.1 | 3.39 | 0.28 | 3.12 | 3.8537 | 3.58 | 2.51 | 3.41 | 2.51 | 0.83 | #N/A
| 1 | 3 | 3.2174 | 3.206 | 2.95 | 3.03 | 3.08 | 3.1 | 3.03 | 3.3896 | 3.36 | 2.93 | 2.94 | 3.02 | 3.2249 | 2.94 | 3.2488 | 3.1 | 3.21 | 0.3 | 3.06 | 4.036 | 3.46 | 2.7 | 3.49 | 2.85 | 0.69 | 16
| 1 | 4 | 3.188 | 3.12 | 3.03 | 3.19 | 3.07 | 3.28 | 3.06 | 3.3594 | 3.44 | 2.98 | 2.65 | 3.03 | 3.212 | 2.96 | 3.1921 | 3.16 | 3.19 | 0.32 | 3.08 | 3.9401 | 3.52 | 2.51 | 3.32 | 2.49 | 0.83 | 3.1436
| 1 | 5 | 2.943 | 3.27 | 3 | 3.04 | 3.06 | 3.31 | 2.91 | 3.4018 | 3.44 | 3 | 3.01 | 3.09 | 3.1588 | 2.97 | 3.1461 | 3.27 | 3.19 | 0.35 | 3.05 | 3.9351 | 3.53 | 2.76 | 3.23 | 2.33 | 0.8 | 30
| 1 | 6 | 3.0991 | 3.171 | 2.93 | 3.13 | 3.12 | 3.19 | 3.17 | 3.4202 | 3.4 | 2.9 | 2.99 | 3.19 | 3.2126 | 3.08 | 3.0846 | 3.07 | 3.13 | 0.44 | 3.07 | 3.9658 | 3.33 | 2.83 | 3.09 | 2.48 | 0.73
| Mean |  | 3.1338 | 3.211 | 2.9933 | 3.07 | 3.0767 | 3.2367 | 3.035 | 3.3608 | 3.3833 | 2.9467 | 2.9367 | 3.095 | 3.2227 | 3.005 | 3.1547 | 3.1583 | 3.225 | 0.31833 | 3.0883 | 3.9353 | 3.485 | 2.6867 | 3.3333 | 2.54 | 0.78833
| Median |  | 3.1758 | 3.2005 | 3.01 | 3.045 | 3.07 | 3.26 | 3.045 | 3.3745 | 3.385 | 2.935 | 2.975 | 3.07 | 3.2165 | 2.99 | 3.1443 | 3.13 | 3.2 | 0.31 | 3.075 | 3.9376 | 3.505 | 2.73 | 3.365 | 2.5 | 0.815
| Std Dev. |  | 0.10174 | 0.066771 | 0.043205 | 0.076158 | 0.022509 | 0.080166 | 0.14124 | 0.053564 | 0.056804 | 0.036697 | 0.14747 | 0.077395 | 0.047889 | 0.058907 | 0.058364 | 0.084242 | 0.088487 | 0.073869 | 0.038687 | 0.064327 | 0.085965 | 0.14404 | 0.15267 | 0.17251 | 0.064005
| Rel.Std.Dev. |  | 3.246% | 2.079% | 1.443% | 2.481% | 0.7316% | 2.477% | 4.654% | 1.594% | 1.679% | 1.245% | 5.022% | 2.501% | 1.486% | 1.96% | 1.85% | 2.667% | 2.744% | 23.21% | 1.253% | 1.635% | 2.467% | 5.361% | 4.58% | 6.792% | 8.119%
| PDM3 |  | -0.3133% | 2.143% | -4.781% | -2.342% | -2.13% | 2.96% | -3.455% | 6.908% | 7.625% | -6.265% | -6.583% | -1.547% | 2.515% | -4.409% | 0.3527% | 0.4681% | 2.589% | -89.87% | -1.759% | 25.18% | 10.86% | -14.54% | 6.035% | -19.2% | -74.92%
| Z-Score (Absolute) |  | 0.2 | 0.6 | 0.53 | 0.13 | 0.1 | 0.73 | 0.31 | 1.38 | 1.5 | 0.77 | 0.82 | 0 | 0.66 | 0.47 | 0.31 | 0.33 | 0.67 | 14.4 | 0.03 | 4.36 | 2.02 | 2.12 | 1.24 | 2.88 | 11.96
| Analytical results for La in OREAS 501b (Certified Value 33 ppm) | [26 empty] | Certified
La | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
1
| 1 | 1 | 37.075 | 36.23 | 32.2 | 36.8 | 29 | 31.3 | 31.5 | 36.31 | 36.8 | 32 | 42 | 40 | 35 | 30 | 34.98 | 35.2 | 31.5 | 27 | 30.9 | 34 | 36.2 | 35.7 | 27 | 22 |  | 1
| 1 | 2 | 37.187 | 36.43 | 30.4 | 34.5 | 27.7 | 30.9 | 32 | 34.12 | 36.4 | 33 | 42 | 30 | 33 | 30 | 33.53 | 32.4 | 30.1 | 27 | 32 | 35.9 | 37.8 | 40.6 | 28 | 19 |  | #N/A
| 1 | 3 | 38.119 | 35.6 | 32.9 | 37.5 | 31.7 | 31.1 | 31.1 | 36.1 | 35.4 | 31 | 43 | 30 | 33 | 30 | 33.16 | 32.7 | 27 | 25 | 30.9 | 37.4 | 38.2 | 32.5 | 29 | 20 |  | 16
| 1 | 4 | 39.481 | 38.47 | 32 | 38.7 | 31.1 | 30.1 | 31 | 34.65 | 38.6 | 34 | 22 | 30 | 34 | 30 | 34.97 | 31.3 | 27.9 | 25 | 32.7 | 37.3 | 37.6 | 31 | 33 | 24 |  | 32.9652
| 1 | 5 | 33.715 | 36.6 | 32.5 | 36.5 | 29.2 | 31 | 31.3 | 34.89 | 37.7 | 33 | 44 | 30 | 33 | 30 | 33 | 32 | 26.9 | 29 | 32.4 | 38.8 | 39.1 | 27.1 | 32 | 22 |  | 31
| 1 | 6 | 33.247 | 35.9 | 30.7 | 39.7 | 31.6 | 29.3 | 32.3 | 35.74 | 37.8 | 33 | 45 | 30 | 35 | 30 | 33.16 | 33.4 | 33.1 | 27 | 27.4 | 36.8 | 39.5 | 31.5 | 30 | 23
| Mean |  | 36.4707 | 36.5383 | 31.7833 | 37.2833 | 30.05 | 30.6167 | 31.5333 | 35.3017 | 37.1167 | 32.6667 | 39.6667 | 31.6667 | 33.8333 | 30 | 33.8 | 32.8333 | 29.4167 | 26.6667 | 31.05 | 36.7 | 38.0667 | 33.0667 | 29.8333 | 21.6667
| Median |  | 37.131 | 36.33 | 32.1 | 37.15 | 30.15 | 30.95 | 31.4 | 35.315 | 37.25 | 33 | 42.5 | 30 | 33.5 | 30 | 33.345 | 32.55 | 29 | 27 | 31.45 | 37.05 | 38 | 32 | 29.5 | 22
| Std Dev. |  | 2.4756 | 1.0129 | 1.0068 | 1.8159 | 1.6477 | 0.76529 | 0.5164 | 0.87598 | 1.1462 | 1.0328 | 8.7331 | 4.0825 | 0.98319 | 0 | 0.92665 | 1.3545 | 2.5631 | 1.5055 | 1.9398 | 1.6248 | 1.1759 | 4.6116 | 2.3166 | 1.8619
| Rel.Std.Dev. |  | 6.788% | 2.772% | 3.168% | 4.871% | 5.483% | 2.5% | 1.638% | 2.481% | 3.088% | 3.162% | 22.02% | 12.89% | 2.906% | 0% | 2.742% | 4.125% | 8.713% | 5.646% | 6.247% | 4.427% | 3.089% | 13.95% | 7.765% | 8.593%
| PDM3 |  | 10.63% | 10.84% | -3.585% | 13.1% | -8.843% | -7.124% | -4.344% | 7.088% | 12.59% | -0.9056% | 20.33% | -3.939% | 2.633% | -8.995% | 2.532% | -0.4% | -10.76% | -19.11% | -5.81% | 11.33% | 15.48% | 0.3078% | -9.501% | -34.27%
| Z-Score (Absolute) |  | 0.92 | 0.94 | 0.24 | 1.12 | 0.67 | 0.53 | 0.3 | 0.63 | 1.08 | 0.02 | 1.71 | 0.27 | 0.27 | 0.68 | 0.26 | 0.02 | 0.82 | 1.51 | 0.42 | 0.98 | 1.32 | 0.08 | 0.72 | 2.74
| Analytical results for Li in OREAS 501b (Certified Value 33.2 ppm) | [26 empty] | Certified
Li | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 1
2
| 1 | 1 | 34.4374 | 33 | 30.5 | 27.5 | 26.6 | 34.6 | 38 | 34.6 | 35 | 32 | 35.6 | 30 | 34.9324 | 33.7 | 34.3 | 7 | 30.3 | 32 | 31 | 33 | 36 | 17 |  |  |  | 1
| 1 | 2 | 32.8808 | 34.2 | 28 | 27.2 | 26 | 35.7 | 38 | 34.6 | 36 | 32 | 36 | 31 | 35.9669 | 33.4 | 38.1 | 7 | 29.8 | 33 | 30 | 34 | 36 | 17 |  |  |  | #N/A
| 1 | 3 | 34.0183 | 34.4 | 29 | 28.4 | 26.1 | 36.9 | 38 | 34.6 | 37 | 31 | 36.9 | 31 | 35.2831 | 33.9 | 35.2 | 5 | 30.5 | 32 | 30 | 33 | 36 | 15 |  |  |  | 16
| 1 | 4 | 33.9691 | 36.1 | 30 | 30.3 | 26.2 | 39.6 | 38 | 35.2 | 37 | 33 | 36.2 | 31 | 36.2779 | 33.8 | 36.2 | 5 | 33 | 33 | 30 | 34 | 39 | 17 |  |  |  | 33.1901
| 1 | 5 | 31.3616 | 33.1 | 31.5 | 28.2 | 25.8 | 38.2 | 38 | 34.6 | 38 | 32 | 36.3 | 30 | 33.7678 | 34.5 | 35.2 | 10 | 30.2 | 33 | 30 | 32 | 37 | 16 |  |  |  | 32
| 1 | 6 | 31.8257 | 34.1 | 30 | 29.9 | 26.8 | 36.9 | 38 | 35.3 | 38 | 32 | 35.1 | 31 | 33.876 | 33.6 | 35 | 8 | 29.4 | 32 | 32 | 32 | 37 | 16
| Mean |  | 33.0821 | 34.15 | 29.8333 | 28.5833 | 26.25 | 36.9833 | 38 | 34.8167 | 36.8333 | 32 | 36.0167 | 30.6667 | 35.0174 | 33.8167 | 35.6667 | 7 | 30.5333 | 32.5 | 30.5 | 33 | 36.8333 | 16.3333
| Median |  | 33.425 | 34.15 | 30 | 28.3 | 26.15 | 36.9 | 38 | 34.6 | 37 | 32 | 36.1 | 31 | 35.1078 | 33.75 | 35.2 | 7 | 30.25 | 32.5 | 30 | 33 | 36.5 | 16.5
| Std Dev. |  | 1.2712 | 1.1221 | 1.2111 | 1.2608 | 0.37815 | 1.7702 | 0 | 0.33714 | 1.169 | 0.63246 | 0.61779 | 0.5164 | 1.0424 | 0.37639 | 1.3382 | 1.8974 | 1.2707 | 0.54772 | 0.83666 | 0.89443 | 1.169 | 0.8165
| Rel.Std.Dev. |  | 3.842% | 3.286% | 4.059% | 4.411% | 1.441% | 4.787% | 0% | 0.9683% | 3.174% | 1.976% | 1.715% | 1.684% | 2.977% | 1.113% | 3.752% | 27.11% | 4.162% | 1.685% | 2.743% | 2.71% | 3.174% | 4.999%
| PDM3 |  | -0.3254% | 2.892% | -10.11% | -13.88% | -20.91% | 11.43% | 14.49% | 4.901% | 10.98% | -3.586% | 8.516% | -7.603% | 5.505% | 1.888% | 7.462% | -78.91% | -8.005% | -2.079% | -8.105% | -0.5729% | 10.98% | -50.79%
| Z-Score (Absolute) |  | 0.01 | 0.29 | 0.84 | 1.16 | 1.77 | 1.03 | 1.29 | 0.46 | 0.99 | 0.27 | 0.78 | 0.62 | 0.52 | 0.2 | 0.69 | 6.8 | 0.65 | 0.14 | 0.66 | 0.01 | 0.99 | 4.36
| Analytical results for Lu in OREAS 501b (Certified Value 0.38 ppm) | [26 empty] | Certified
Lu | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab |  |  |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 16; 06: 23; 07: 24; 08: 25 |  |  |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS |  |  |  | 2
3
| 1 | 1 | 01: 0.3973; 02: 0.374; 03: 0.34; 04: 0.345; 05: 0.40353; 06: 0.4; 07: 0.36; 08: 0.38 |  |  |  | 1
| 1 | 2 | 01: 0.3367; 02: 0.375; 03: 0.32; 04: 0.353; 05: 0.40038; 06: 0.4; 07: 0.37; 08: 0.43 |  |  |  | #N/A
| 1 | 3 | 01: 0.422; 02: 0.364; 03: 0.32; 04: 0.371; 05: 0.40853; 06: 0.4; 07: 0.36; 08: 0.39 |  |  |  | 16
| 1 | 4 | 01: 0.3351; 02: 0.384; 03: 0.32; 04: 0.372; 05: 0.40117; 06: 0.4; 07: 0.36; 08: 0.35 |  |  |  | 0.37503
| 1 | 5 | 01: 0.4193; 02: 0.362; 03: 0.32; 04: 0.367; 05: 0.40169; 06: 0.4; 07: 0.38; 08: 0.44 |  |  |  | 33
| 1 | 6 | 01: 0.3304; 02: 0.376; 03: 0.32; 04: 0.362; 05: 0.40042; 06: 0.3; 07: 0.38; 08: 0.42
| Mean |  | 01: 0.37347; 02: 0.3725; 03: 0.32333; 04: 0.36167; 05: 0.40262; 06: 0.38333; 07: 0.36833; 08: 0.40167
| Median |  | 01: 0.367; 02: 0.3745; 03: 0.32; 04: 0.3645; 05: 0.40143; 06: 0.4; 07: 0.365; 08: 0.405
| Std Dev. |  | 01: 0.044052; 02: 0.0081915; 03: 0.008165; 04: 0.010727; 05: 0.0031161; 06: 0.040825; 07: 0.0098319; 08: 0.034303
| Rel.Std.Dev. |  | 01: 11.8%; 02: 2.199%; 03: 2.525%; 04: 2.966%; 05: 0.774%; 06: 10.65%; 07: 2.669%; 08: 8.54%
| PDM3 |  | 01: -0.4172%; 02: -0.675%; 03: -13.79%; 04: -3.564%; 05: 7.356%; 06: 2.214%; 07: -1.786%; 08: 7.102%
| Z-Score (Absolute) |  | 01: 0.03; 02: 0.03; 03: 3.09; 04: 0.7; 05: 1.84; 06: 0.64; 07: 0.29; 08: 1.79
| Analytical results for Mg in OREAS 501b (Certified Value 1.48 wt.%) | [26 empty] | Certified
Mg | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 2
3
| 1 | 1 | 1.541 | 1.519 | 1.61 | 1.45 | 1.39 | 1.49 | 1.49 | 1.4461 | 1.47 | 1.53 | 1.74 | 1.53 | 1.4531 | 1.44 | 1.483 | 1.52 | 1.47 | 1.28 | 1.46 | 1.5957 | 1.64 | 1.54 | 1.54 | 1.37 | 1.21 | 1
| 1 | 2 | 1.5197 | 1.552 | 1.59 | 1.43 | 1.39 | 1.49 | 1.49 | 1.4445 | 1.5 | 1.54 | 1.7 | 1.43 | 1.4985 | 1.45 | 1.489 | 1.44 | 1.55 | 1.24 | 1.46 | 1.5401 | 1.61 | 1.51 | 1.51 | 1.44 | 1.19 | #N/A
| 1 | 3 | 1.5716 | 1.551 | 1.59 | 1.43 | 1.39 | 1.47 | 1.49 | 1.4761 | 1.51 | 1.51 | 1.73 | 1.45 | 1.4563 | 1.42 | 1.515 | 1.44 | 1.45 | 1.23 | 1.42 | 1.5962 | 1.68 | 1.52 | 1.55 | 1.37 | 1.11 | 16
| 1 | 4 | 1.5555 | 1.55 | 1.59 | 1.48 | 1.39 | 1.46 | 1.49 | 1.4714 | 1.52 | 1.56 | 1.43 | 1.46 | 1.4616 | 1.42 | 1.502 | 1.43 | 1.43 | 1.24 | 1.4 | 1.5662 | 1.62 | 1.52 | 1.48 | 1.41 | 1.27 | 1.4821
| 1 | 5 | 1.4204 | 1.536 | 1.61 | 1.44 | 1.4 | 1.49 | 1.48 | 1.4765 | 1.53 | 1.51 | 1.75 | 1.44 | 1.4364 | 1.43 | 1.486 | 1.49 | 1.43 | 1.36 | 1.47 | 1.5826 | 1.65 | 1.53 | 1.42 | 1.44 | 1.22 | 34
| 1 | 6 | 1.4975 | 1.52 | 1.59 | 1.52 | 1.41 | 1.46 | 1.47 | 1.4984 | 1.54 | 1.52 | 1.71 | 1.44 | 1.4548 | 1.48 | 1.473 | 1.42 | 1.43 | 1.29 | 1.4 | 1.5815 | 1.63 | 1.55 | 1.38 | 1.41 | 1.22
| Mean |  | 1.5176 | 1.538 | 1.5967 | 1.4583 | 1.395 | 1.4767 | 1.485 | 1.4688 | 1.5117 | 1.5283 | 1.6767 | 1.4583 | 1.4601 | 1.44 | 1.4913 | 1.4567 | 1.46 | 1.2733 | 1.435 | 1.5771 | 1.6383 | 1.5283 | 1.48 | 1.4067 | 1.2033
| Median |  | 1.5304 | 1.543 | 1.59 | 1.445 | 1.39 | 1.48 | 1.49 | 1.4737 | 1.515 | 1.525 | 1.72 | 1.445 | 1.4555 | 1.435 | 1.4875 | 1.44 | 1.44 | 1.26 | 1.44 | 1.582 | 1.635 | 1.525 | 1.495 | 1.41 | 1.215
| Std Dev. |  | 0.054325 | 0.015479 | 0.010328 | 0.035449 | 0.0083666 | 0.015055 | 0.0083666 | 0.0205 | 0.024833 | 0.019408 | 0.12226 | 0.03656 | 0.02064 | 0.022804 | 0.014922 | 0.039328 | 0.046904 | 0.048854 | 0.032094 | 0.021205 | 0.024833 | 0.01472 | 0.067823 | 0.031411 | 0.052789
| Rel.Std.Dev. |  | 3.58% | 1.006% | 0.6468% | 2.431% | 0.5998% | 1.02% | 0.5634% | 1.396% | 1.643% | 1.27% | 7.292% | 2.507% | 1.414% | 1.584% | 1.001% | 2.7% | 3.213% | 3.837% | 2.236% | 1.345% | 1.516% | 0.9631% | 4.583% | 2.233% | 4.387%
| PDM3 |  | 2.394% | 3.769% | 7.727% | -1.606% | -5.879% | -0.3693% | 0.1929% | -0.8978% | 1.992% | 3.117% | 13.12% | -1.606% | -1.486% | -2.843% | 0.6202% | -1.719% | -1.494% | -14.09% | -3.181% | 6.404% | 10.54% | 3.117% | -0.1444% | -5.092% | -18.81%
| Z-Score (Absolute) |  | 0.67 | 1.01 | 1.98 | 0.3 | 1.34 | 0 | 0.14 | 0.13 | 0.58 | 0.85 | 3.29 | 0.3 | 0.27 | 0.6 | 0.24 | 0.33 | 0.27 | 3.35 | 0.69 | 1.65 | 2.66 | 0.85 | 0.05 | 1.15 | 4.5
| Analytical results for Mn in OREAS 501b (Certified Value 546 ppm) | [26 empty] | Certified
Mn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 561.031 | 577 | 580 | 538 | 546 | 572 | 514 | 561 | 514 | 552 | 634 | 562 | 555 | 540 | 518.3149 | 555 | 553 | 450 | 547 | 558 | 553 | 578 | 500 | 488 | 440 | 1
| 1 | 2 | 549.464 | 589 | 580 | 527 | 545 | 573 | 510 | 557 | 526 | 549 | 618 | 538 | 566 | 540 | 520.3455 | 527 | 581 | 428 | 541 | 559 | 548 | 559 | 500 | 528 | 440 | #N/A
| 1 | 3 | 565.028 | 590 | 584 | 537 | 546 | 558 | 515 | 570 | 521 | 530 | 629 | 540 | 551 | 527 | 533.2969 | 529 | 552 | 424 | 515 | 561 | 540 | 576 | 500 | 492 | 400 | 16
| 1 | 4 | 560.07 | 590 | 590 | 563 | 544 | 545 | 512 | 567 | 526 | 562 | 546 | 540 | 580 | 528 | 531.5011 | 543 | 550 | 429 | 510 | 574 | 545 | 566 | 500 | 589 | 460 | 546.0468
| 1 | 5 | 509.435 | 583 | 580 | 538 | 545 | 583 | 509 | 569 | 528 | 538 | 634 | 543 | 581 | 531 | 522.6845 | 554 | 543 | 478 | 523 | 572 | 561 | 565 | 500 | 510 | 440 | 35
| 1 | 6 | 542.678 | 577 | 584 | 552 | 555 | 549 | 515 | 580 | 538 | 565 | 622 | 557 | 572 | 548 | 521.7115 | 522 | 533 | 461 | 528 | 567 | 542 | 587 | 500 | 509 | 440
| Mean |  | 547.951 | 584.3333 | 583 | 542.5 | 546.8333 | 563.3333 | 512.5 | 567.3333 | 525.5 | 549.3333 | 613.8333 | 546.6667 | 567.5 | 535.6667 | 524.6424 | 538.3333 | 552 | 445 | 527.3333 | 565.1667 | 548.1667 | 571.8333 | 500 | 519.3333 | 436.6667
| Median |  | 554.767 | 586 | 582 | 538 | 545.5 | 565 | 513 | 568 | 526 | 550.5 | 625.5 | 541.5 | 569 | 535.5 | 522.198 | 536 | 551 | 439.5 | 525.5 | 564 | 546.5 | 571 | 500 | 509.5 | 440
| Std Dev. |  | 20.6066 | 6.2503 | 3.9497 | 12.818 | 4.0702 | 15.0022 | 2.5884 | 7.9666 | 7.9436 | 13.5302 | 33.8492 | 10.1915 | 12.5658 | 8.3106 | 6.2107 | 14.3341 | 16.0499 | 21.7071 | 14.4591 | 6.8532 | 7.7825 | 10.3037 | 0 | 37.0171 | 19.6638
| Rel.Std.Dev. |  | 3.761% | 1.07% | 0.6775% | 2.363% | 0.7443% | 2.663% | 0.5051% | 1.404% | 1.512% | 2.463% | 5.514% | 1.864% | 2.214% | 1.551% | 1.184% | 2.663% | 2.908% | 4.878% | 2.742% | 1.213% | 1.42% | 1.802% | 0% | 7.128% | 4.503%
| PDM3 |  | 0.3487% | 7.012% | 6.767% | -0.6495% | 0.144% | 3.166% | -6.144% | 3.898% | -3.763% | 0.6019% | 12.41% | 0.1135% | 3.929% | -1.901% | -3.92% | -1.413% | 1.09% | -18.51% | -3.427% | 3.502% | 0.3882% | 4.722% | -8.433% | -4.892% | -20.03%
| Z-Score (Absolute) |  | 0.04 | 1.23 | 1.19 | 0.14 | 0 | 0.54 | 1.13 | 0.67 | 0.7 | 0.08 | 2.2 | 0.01 | 0.68 | 0.37 | 0.73 | 0.28 | 0.17 | 3.35 | 0.64 | 0.6 | 0.04 | 0.82 | 1.54 | 0.9 | 3.62
| Analytical results for Mo in OREAS 501b (Certified Value 99 ppm) | [26 empty] | Certified
Mo | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 102.427 | 105.6 | 105 | 102 | 93.6 | 102.7 | 86 | 100.7 | 80 | 107 | 113 | 101 | 101 | 97 | 100.771 | 103.5 | 103 | 92 | 101 | 83 | 97 | 116 | 88.08 | 92 | 67 | 1
| 1 | 2 | 101.667 | 104.7 | 95.5 | 99.4 | 91.8 | 102.7 | 84 | 105.2 | 82 | 107 | 109 | 98 | 101 | 98 | 100.233 | 99.9 | 110.5 | 86 | 103 | 78 | 99 | 107 | 89.76 | 95 | 65 | #N/A
| 1 | 3 | 103.765 | 103.3 | 99 | 99.7 | 93.2 | 98.6 | 85 | 108.3 | 80 | 104 | 110 | 101 | 100 | 95 | 100.447 | 98.7 | 103 | 87 | 99.6 | 79 | 96 | 109 | 86.5 | 94 | 61 | 16
| 1 | 4 | 101.044 | 107 | 100 | 106.5 | 93.3 | 101 | 85 | 105.6 | 86 | 103 | 99 | 100 | 99 | 96 | 101.739 | 99 | 106.5 | 88 | 102.5 | 80 | 98 | 108 | 88.07 | 97 | 68 | 98.9612
| 1 | 5 | 92.9091 | 106.9 | 100 | 101 | 96.5 | 101 | 83 | 106.9 | 84 | 102 | 112 | 100 | 96 | 98 | 101.457 | 103.5 | 105 | 95 | 99.3 | 79 | 100 | 115 | 83.38 | 92 | 69 | 36
| 1 | 6 | 102.835 | 107.7 | 99.5 | 105.5 | 94.1 | 98.2 | 84 | 106.8 | 87 | 104 | 109 | 99 | 100 | 101 | 100.828 | 97.1 | 103 | 93 | 98 | 81 | 97 | 108 | 84.63 | 94 | 68
| Mean |  | 100.7745 | 105.8667 | 99.8333 | 102.35 | 93.75 | 100.7 | 84.5 | 105.5833 | 83.1667 | 104.5 | 108.6667 | 99.8333 | 99.5 | 97.5 | 100.9125 | 100.2833 | 105.1667 | 90.1667 | 100.5667 | 80 | 97.8333 | 110.5 | 86.7367 | 94 | 66.3333
| Median |  | 102.047 | 106.25 | 99.75 | 101.5 | 93.45 | 101 | 84.5 | 106.2 | 83 | 104 | 109.5 | 100 | 100 | 97.5 | 100.7995 | 99.45 | 104 | 90 | 100.3 | 79.5 | 97.5 | 108.5 | 87.285 | 94 | 67.5
| Std Dev. |  | 3.9662 | 1.6573 | 3.0441 | 2.9938 | 1.5502 | 1.9411 | 1.0488 | 2.6301 | 2.9944 | 2.0736 | 5.0067 | 1.169 | 1.8708 | 2.0736 | 0.58054 | 2.6506 | 2.9777 | 3.656 | 1.948 | 1.7889 | 1.472 | 3.937 | 2.3868 | 1.8974 | 2.9439
| Rel.Std.Dev. |  | 3.936% | 1.565% | 3.049% | 2.925% | 1.654% | 1.928% | 1.241% | 2.491% | 3.601% | 1.984% | 4.607% | 1.171% | 1.88% | 2.127% | 0.5753% | 2.643% | 2.831% | 4.055% | 1.937% | 2.236% | 1.505% | 3.563% | 2.752% | 2.018% | 4.438%
| PDM3 |  | 1.832% | 6.978% | 0.8813% | 3.424% | -5.266% | 1.757% | -14.61% | 6.692% | -15.96% | 5.597% | 9.807% | 0.8813% | 0.5445% | -1.477% | 1.972% | 1.336% | 6.271% | -8.887% | 1.622% | -19.16% | -1.14% | 11.66% | -12.35% | -5.013% | -32.97%
| Z-Score (Absolute) |  | 0.12 | 0.76 | 0 | 0.32 | 0.77 | 0.11 | 1.94 | 0.73 | 2.11 | 0.59 | 1.12 | 0 | 0.04 | 0.3 | 0.14 | 0.06 | 0.67 | 1.22 | 0.09 | 2.51 | 0.25 | 1.35 | 1.66 | 0.74 | 4.24
| Analytical results for Na in OREAS 501b (Certified Value 2.08 wt.%) | [26 empty] | Certified
Na | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 20 | 22 | 23 | 24 | 25 | 26 | 27 |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  | 2
3
| 1 | 1 | 2.1785 | 2.003 | 2.04 | 2.02 | 2.05 | 2.206 | 2.13 | 2.2022 | 2.15 | 1.95 | 2.36 | 2.12 | 2.0058 | 1.98 | 2.0892 | 2.1 | 2.11 | 2.05 | 2.2565 | 2.19 | 1.94 | 2.18 | 1.65 | 1.63 |  | 1
| 1 | 2 | 2.0895 | 2.04 | 2.05 | 1.98 | 2.06 | 2.192 | 2.14 | 2.1968 | 2.18 | 1.97 | 2.31 | 2.05 | 2.0452 | 2.01 | 2.0982 | 2.03 | 2.2 | 2.04 | 2.156 | 2.28 | 1.85 | 2.11 | 1.39 | 1.63 |  | #N/A
| 1 | 3 | 2.1991 | 2.047 | 2.02 | 1.96 | 2.06 | 2.161 | 2.08 | 2.2741 | 2.22 | 1.91 | 2.33 | 2.01 | 2.0017 | 1.95 | 2.147 | 1.99 | 2.1 | 2.01 | 2.2487 | 2.33 | 1.86 | 2.17 | 1.49 | 1.53 |  | 16
| 1 | 4 | 2.1625 | 2.046 | 2.06 | 2.07 | 2.04 | 2.163 | 2.14 | 2.2586 | 2.25 | 1.96 | 1.92 | 2.02 | 1.9977 | 1.92 | 2.1208 | 2 | 2.1 | 2.03 | 2.2244 | 2.23 | 1.82 | 2.06 | 1.29 | 1.69 |  | 2.0848
| 1 | 5 | 1.9141 | 2.024 | 2.03 | 1.99 | 2.04 | 2.189 | 2.11 | 2.2909 | 2.31 | 1.9 | 2.37 | 2.06 | 1.9729 | 1.96 | 2.1213 | 2.1 | 2.08 | 2.07 | 2.2287 | 2.25 | 1.84 | 2.01 | 1.3 | 1.63 |  | 37
| 1 | 6 | 2.1577 | 2.02 | 2 | 2.11 | 2.06 | 2.148 | 2.17 | 2.2837 | 2.38 | 1.92 | 2.33 | 2.13 | 2.0051 | 2.03 | 2.0306 | 1.98 | 2.04 | 2 | 2.2433 | 2.21 | 1.95 | 1.97 | 1.3 | 1.6
| Mean |  | 2.1169 | 2.03 | 2.0333 | 2.0217 | 2.0517 | 2.1765 | 2.1283 | 2.2511 | 2.2483 | 1.935 | 2.27 | 2.065 | 2.0047 | 1.975 | 2.1012 | 2.0333 | 2.105 | 2.0333 | 2.2263 | 2.2483 | 1.8767 | 2.0833 | 1.4033 | 1.6183
| Median |  | 2.1601 | 2.032 | 2.035 | 2.005 | 2.055 | 2.176 | 2.135 | 2.2664 | 2.235 | 1.935 | 2.33 | 2.055 | 2.0034 | 1.97 | 2.1095 | 2.015 | 2.1 | 2.035 | 2.236 | 2.24 | 1.855 | 2.085 | 1.345 | 1.63
| Std Dev. |  | 0.10599 | 0.017378 | 0.021602 | 0.057764 | 0.0098319 | 0.022367 | 0.030605 | 0.041405 | 0.085186 | 0.02881 | 0.17286 | 0.0501 | 0.023285 | 0.040373 | 0.04006 | 0.054283 | 0.05282 | 0.02582 | 0.036481 | 0.050761 | 0.05465 | 0.085245 | 0.14334 | 0.052313
| Rel.Std.Dev. |  | 5.007% | 0.8561% | 1.062% | 2.857% | 0.4792% | 1.028% | 1.438% | 1.839% | 3.789% | 1.489% | 7.615% | 2.426% | 1.161% | 2.044% | 1.907% | 2.67% | 2.509% | 1.27% | 1.639% | 2.258% | 2.912% | 4.092% | 10.21% | 3.233%
| PDM3 |  | 1.54% | -2.628% | -2.468% | -3.028% | -1.589% | 4.399% | 2.089% | 7.975% | 7.845% | -7.185% | 8.884% | -0.9493% | -3.84% | -5.266% | 0.7868% | -2.468% | 0.9694% | -2.468% | 6.786% | 7.845% | -9.983% | -0.06987% | -32.69% | -22.37%
| Z-Score (Absolute) |  | 0.61 | 0.3 | 0.26 | 0.38 | 0.07 | 1.24 | 0.73 | 2.02 | 1.99 | 1.29 | 2.22 | 0.07 | 0.56 | 0.87 | 0.45 | 0.26 | 0.49 | 0.26 | 1.76 | 1.99 | 1.91 | 0.26 | 6.87 | 4.62
| Analytical results for Nb in OREAS 501b (Certified Value 18 ppm) | [26 empty] | Certified
Nb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 10 | 11 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 | [5 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 1
1
| 1 | 1 | 18.385 | 17.16 | 17.5 | 17.8 | 15.6 | 17.9 | 19.9 | 16 | 19 | 18.5 | 24.83 | 17 | 19.9 | 14 | 17.2 | 19 | 19 | 22.3 | 33 | 15 | [5 empty] | 1
| 1 | 2 | 17.94 | 17.47 | 16.5 | 17.4 | 15.2 | 18.5 | 20.14 | 16 | 18 | 19 | 21.81 | 16.7 | 21.5 | 11 | 18.1 | 18 | 18.5 | 19.3 | 32 | 14 | [5 empty] | 10
| 1 | 3 | 17.033 | 17.87 | 16.5 | 18 | 15.9 | 18.5 | 20.16 | 15 | 18 | 19 | 21.59 | 16.6 | 19.7 | 12 | 17.8 | 19 | 18.8 | 21.1 | 26 | 13 | [5 empty] | 16
| 1 | 4 | 18.096 | 18.38 | 17.5 | 18.6 | 15.3 | 17.3 | 20.2 | 16 | 16 | 18.8 | 21.46 | 17.1 | 20.4 | 13 | 18.1 | 16 | 18.6 | 20 | 32 | 15 | [5 empty] | 18.0216
| 1 | 5 | 16.037 | 18.28 | 17.5 | 17.8 | 16.1 | 18.5 | 19.88 | 16 | 18 | 18.7 | 21.25 | 17.1 | 19.7 | 14 | 18.1 | 18 | 19.3 | 27.5 | 32 | 14 | [5 empty] | 38
| 1 | 6 | 15.736 | 17.9 | 17 | 18.4 | 16.2 | 17.9 | 19.93 | 16 | 18 | 18.8 | 21.05 | 16.4 | 19.3 | 11 | 17.9 | 16 | 19.9 | 20.5 | 32 | 14
| Mean |  | 17.2045 | 17.8433 | 17.0833 | 18 | 15.7167 | 18.1 | 20.035 | 15.8333 | 17.8333 | 18.8 | 21.9983 | 16.8167 | 20.0833 | 12.5 | 17.8667 | 17.6667 | 19.0167 | 21.7833 | 31.1667 | 14.1667
| Median |  | 17.4865 | 17.885 | 17.25 | 17.9 | 15.75 | 18.2 | 20.035 | 16 | 18 | 18.8 | 21.525 | 16.85 | 19.8 | 12.5 | 18 | 18 | 18.9 | 20.8 | 32 | 14
| Std Dev. |  | 1.1206 | 0.46668 | 0.4916 | 0.43818 | 0.41673 | 0.4899 | 0.14639 | 0.40825 | 0.98319 | 0.18974 | 1.4121 | 0.29269 | 0.78081 | 1.3784 | 0.35024 | 1.3663 | 0.51929 | 2.9802 | 2.5626 | 0.75277
| Rel.Std.Dev. |  | 6.513% | 2.615% | 2.878% | 2.434% | 2.652% | 2.707% | 0.7307% | 2.578% | 5.513% | 1.009% | 6.419% | 1.74% | 3.888% | 11.03% | 1.96% | 7.734% | 2.731% | 13.68% | 8.222% | 5.314%
| PDM3 |  | -4.534% | -0.989% | -5.206% | -0.1197% | -12.79% | 0.4352% | 11.17% | -12.14% | -1.044% | 4.319% | 22.07% | -6.686% | 11.44% | -30.64% | -0.8595% | -1.969% | 5.522% | 20.87% | 72.94% | -21.39%
| Z-Score (Absolute) |  | 0.4 | 0.01 | 0.47 | 0.09 | 1.31 | 0.15 | 1.34 | 1.24 | 0.01 | 0.58 | 2.54 | 0.64 | 1.37 | 3.28 | 0.01 | 0.12 | 0.71 | 2.41 | 8.16 | 2.26
| Analytical results for Nd in OREAS 501b (Certified Value 30.6 ppm) | [26 empty] | Certified
Nd | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [5 empty] | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | [5 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | [5 empty] | 1
2
| 1 | 1 | 01: 29.643; 02: 30.81; 03: 26.2; 04: 30.89; 05: 28.5; 06: 28.4 | [5 empty] | 1
| 1 | 2 | 01: 30.48; 02: 32.47; 03: 25.1; 04: 29.64; 05: 30.4; 06: 28.3 | [5 empty] | #N/A
| 1 | 3 | 01: 32.637; 02: 31.48; 03: 26.2; 04: 30.99; 05: 31.4; 06: 29.6 | [5 empty] | 16
| 1 | 4 | 01: 32.126; 02: 32.16; 03: 26.3; 04: 29.78; 05: 31.3; 06: 28.7 | [5 empty] | 30.5668
| 1 | 5 | 01: 27.985; 02: 31.39; 03: 26.7; 04: 30.17; 05: 32; 06: 29.5 | [5 empty] | 39
| 1 | 6 | 01: 29.312; 02: 31.97; 03: 25.2; 04: 30.89; 05: 31.3; 06: 30
| Mean |  | 01: 30.3638; 02: 31.7133; 03: 25.95; 04: 30.3933; 05: 30.8167; 06: 29.0833
| Median |  | 01: 30.0615; 02: 31.725; 03: 26.2; 04: 30.53; 05: 31.3; 06: 29.1
| Std Dev. |  | 01: 1.7647; 02: 0.60215; 03: 0.6473; 04: 0.60711; 05: 1.2449; 06: 0.70828
| Rel.Std.Dev. |  | 01: 5.812%; 02: 1.899%; 03: 2.494%; 04: 1.998%; 05: 4.04%; 06: 2.435%
| PDM3 |  | 01: -0.6639%; 02: 3.751%; 03: -15.1%; 04: -0.5674%; 05: 0.8176%; 06: -4.853%
| Z-Score (Absolute) |  | 01: 0.01; 02: 1.04; 03: 3.45; 04: 0.01; 05: 0.34; 06: 1.01
| Analytical results for Ni in OREAS 501b (Certified Value 41.5 ppm) | [26 empty] | Certified
Ni | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 44.418 | 44 | 44 | 40.6 | 37.6 | 48.3 | 39 | 39.5 | 42 | 46 | 49 | 42.8 | 41 | 40 | 37.58 | 40.5 | 45.6 | 37 | 39.7 | 30 | 41.6 | 42.4 | 51.1 | 37 | 36 | 1
| 1 | 2 | 42.594 | 46 | 42 | 40.3 | 35.4 | 41.5 | 39 | 38.6 | 45 | 44 | 51 | 41.3 | 40 | 38 | 36.94 | 39.7 | 47.4 | 37 | 41.6 | 29 | 42.2 | 42.8 | 46.6 | 37 | 35 | #N/A
| 1 | 3 | 41.988 | 47 | 46 | 40.5 | 37.5 | 42 | 40 | 39.5 | 44 | 44 | 49 | 43 | 40 | 40 | 39.28 | 40.1 | 43.4 | 36 | 38.9 | 28 | 41.3 | 42.6 | 45.7 | 36 | 29 | 16
| 1 | 4 | 44.827 | 45 | 44 | 42.7 | 37.3 | 42.9 | 38 | 40.8 | 46 | 42 | 44 | 43.5 | 39 | 35 | 41.22 | 39.4 | 61.4 | 37 | 41.7 | 28 | 43.8 | 40.8 | 41 | 38 | 34 | 41.4972
| 1 | 5 | 40.194 | 43 | 44 | 40.1 | 36.6 | 44.7 | 38 | 41.6 | 43 | 44 | 48 | 42.8 | 41 | 39 | 40.06 | 42.2 | 71.4 | 39 | 42.2 | 28 | 44 | 43.3 | 49.1 | 39 | 31 | 40
| 1 | 6 | 46.546 | 45 | 42 | 42.4 | 37.9 | 41.8 | 34 | 39.8 | 45 | 45 | 47 | 43.5 | 40 | 39 | 41.97 | 38 | 43 | 40 | 40.1 | 29 | 41.4 | 43.8 | 45.8 | 37 | 32
| Mean |  | 43.4278 | 45 | 43.6667 | 41.1 | 37.05 | 43.5333 | 38 | 39.9667 | 44.1667 | 44.1667 | 48 | 42.8167 | 40.1667 | 38.5 | 39.5083 | 39.9833 | 52.0333 | 37.6667 | 40.7 | 28.6667 | 42.3833 | 42.6167 | 46.55 | 37.3333 | 32.8333
| Median |  | 43.506 | 45 | 44 | 40.55 | 37.4 | 42.45 | 38.5 | 39.65 | 44.5 | 44 | 48.5 | 42.9 | 40 | 39 | 39.67 | 39.9 | 46.5 | 37 | 40.85 | 28.5 | 41.9 | 42.7 | 46.2 | 37 | 33
| Std Dev. |  | 2.2754 | 1.4142 | 1.5055 | 1.1402 | 0.91815 | 2.6067 | 2.0976 | 1.0671 | 1.472 | 1.3292 | 2.3664 | 0.8085 | 0.75277 | 1.8708 | 1.983 | 1.3819 | 11.6773 | 1.5055 | 1.3161 | 0.8165 | 1.2172 | 1.0245 | 3.4437 | 1.0328 | 2.6394
| Rel.Std.Dev. |  | 5.239% | 3.143% | 3.448% | 2.774% | 2.478% | 5.988% | 5.52% | 2.67% | 3.333% | 3.009% | 4.93% | 1.888% | 1.874% | 4.859% | 5.019% | 3.456% | 22.44% | 3.997% | 3.234% | 2.848% | 2.872% | 2.404% | 7.398% | 2.766% | 8.039%
| PDM3 |  | 4.652% | 8.441% | 5.228% | -0.9572% | -10.72% | 4.907% | -8.428% | -3.688% | 6.433% | 6.433% | 15.67% | 3.18% | -3.206% | -7.223% | -4.793% | -3.648% | 25.39% | -9.231% | -1.921% | -30.92% | 2.135% | 2.698% | 12.18% | -10.03% | -20.88%
| Z-Score (Absolute) |  | 0.6 | 1.01 | 0.67 | 0 | 1.05 | 0.63 | 0.8 | 0.29 | 0.8 | 0.8 | 1.79 | 0.45 | 0.24 | 0.67 | 0.41 | 0.29 | 2.84 | 0.89 | 0.1 | 3.22 | 0.33 | 0.39 | 1.41 | 0.98 | 2.14
| Analytical results for P in OREAS 501b (Certified Value 0.102 wt.%) | [26 empty] | Certified
P | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 24 | 25 | 26 | 27 |  |  |  |  | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  |  | 3
3
| 1 | 1 | 0.099508 | 0.099 | 0.095 | 0.101 | 0.106 | 0.106 | 0.11 | 0.1029 | 0.1 | 0.116 | 0.109 | 0.1042 | 0.1 | 0.107 | 0.104 | 0.083 | 0.108 | 0.0972 | 0.0976 | 0.19 | 0.088 |  |  |  |  | 1
| 1 | 2 | 0.099721 | 0.099 | 0.1 | 0.1 | 0.104 | 0.104 | 0.11 | 0.096 | 0.1 | 0.112 | 0.103 | 0.1028 | 0.099 | 0.102 | 0.108 | 0.0789 | 0.106 | 0.0981 | 0.1005 | 0.16 | 0.088 |  |  |  |  | #N/A
| 1 | 3 | 0.099688 | 0.099 | 0.105 | 0.099 | 0.105 | 0.104 | 0.11 | 0.1045 | 0.09 | 0.114 | 0.107 | 0.1006 | 0.096 | 0.101 | 0.103 | 0.0782 | 0.102 | 0.0982 | 0.0988 | 0.12 | 0.082 |  |  |  |  | 16
| 1 | 4 | 0.1003 | 0.101 | 0.095 | 0.103 | 0.108 | 0.109 | 0.11 | 0.1044 | 0.1 | 0.099 | 0.107 | 0.1012 | 0.099 | 0.1 | 0.102 | 0.0792 | 0.102 | 0.0964 | 0.1002 | 0.1 | 0.092 |  |  |  |  | 0.10246
| 1 | 5 | 0.090304 | 0.1 | 0.105 | 0.1 | 0.106 | 0.11 | 0.11 | 0.1091 | 0.1 | 0.116 | 0.104 | 0.0975 | 0.096 | 0.106 | 0.103 | 0.0889 | 0.106 | 0.0976 | 0.0971 | 0.15 | 0.088 |  |  |  |  | 41
| 1 | 6 | 0.098474 | 0.101 | 0.105 | 0.106 | 0.108 | 0.104 | 0.1 | 0.1045 | 0.1 | 0.113 | 0.104 | 0.1012 | 0.1 | 0.099 | 0.103 | 0.0847 | 0.102 | 0.1022 | 0.0973 | 0.2 | 0.09
| Mean |  | 0.097999 | 0.099833 | 0.10083 | 0.1015 | 0.10617 | 0.10617 | 0.10833 | 0.10357 | 0.098333 | 0.11167 | 0.10567 | 0.10125 | 0.098333 | 0.1025 | 0.10383 | 0.08215 | 0.10433 | 0.098283 | 0.098583 | 0.15333 | 0.088
| Median |  | 0.099598 | 0.0995 | 0.1025 | 0.1005 | 0.106 | 0.105 | 0.11 | 0.10445 | 0.1 | 0.1135 | 0.1055 | 0.1012 | 0.099 | 0.1015 | 0.103 | 0.0811 | 0.104 | 0.09785 | 0.0982 | 0.155 | 0.088
| Std Dev. |  | 0.0038162 | 0.00098319 | 0.004916 | 0.0025884 | 0.0016021 | 0.0027142 | 0.0040825 | 0.0042604 | 0.0040825 | 0.0064083 | 0.0023381 | 0.0022625 | 0.0018619 | 0.0032711 | 0.002137 | 0.004187 | 0.0026583 | 0.0020282 | 0.0014932 | 0.038816 | 0.0033466
| Rel.Std.Dev. |  | 3.894% | 0.9848% | 4.875% | 2.55% | 1.509% | 2.557% | 3.768% | 4.114% | 4.152% | 5.739% | 2.213% | 2.235% | 1.893% | 3.191% | 2.058% | 5.097% | 2.548% | 2.064% | 1.515% | 25.31% | 3.803%
| PDM3 |  | -4.359% | -2.568% | -1.592% | -0.9416% | 3.613% | 3.613% | 5.727% | 1.075% | -4.032% | 8.981% | 3.125% | -1.186% | -4.032% | 0.03438% | 1.336% | -19.83% | 1.824% | -4.081% | -3.788% | 49.64% | -14.12%
| Z-Score (Absolute) |  | 0.75 | 0.35 | 0.14 | 0 | 0.99 | 0.99 | 1.46 | 0.44 | 0.67 | 2.16 | 0.89 | 0.05 | 0.67 | 0.21 | 0.5 | 4.12 | 0.6 | 0.68 | 0.62 | 11.04 | 2.87
| Analytical results for Pb in OREAS 501b (Certified Value 23 ppm) | [26 empty] | Certified
Pb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 1
2
| 1 | 1 | 24.171 | 24.9 | 23 | 24.7 | 19.5 | 24.7 | 24 | 24.4 | 20 | 9 | 24 | 22.3 | 24 | 34 | 26.0725 | 23 | 24 | 14 | 26.5 | < 10 | 25 | 22.2 | 25.8 | 21 | 24 | 1
| 1 | 2 | 23.665 | 24.6 | 21 | 22.4 | 23.1 | 25.2 | 21 | 21.2 | 30 | 7 | 25 | 22.6 | 22 | 28 | 27.1114 | 22.6 | 24.6 | 10 | 23.6 | < 10 | 27 | 22.4 | 25.2 | 25 | 25 | #N/A
| 1 | 3 | 23.198 | 24.5 | 21 | 22 | 19.9 | 24.8 | 21 | 22.8 | 30 | 6 | 22 | 23 | 22 | 22 | 25.9157 | 22.8 | 23.1 | 13 | 24.3 | < 10 | 23 | 21.5 | 23.2 | 18 | 24 | 16
| 1 | 4 | 23.978 | 25.3 | 22 | 23.6 | 20.2 | 25.9 | 22 | 20.6 | 30 | 10 | 17 | 24.5 | 24 | 23 | 25.2339 | 21.9 | 23.7 | 6 | 22.8 | < 10 | 24 | 21.4 | 25.7 | 18 | 26 | 23.0455
| 1 | 5 | 21.913 | 25.2 | 24 | 22.5 | 19.5 | 26.8 | 21 | 21 | 30 | 6 | 22 | 22.5 | 21 | 27 | 22.0103 | 23.1 | 23.4 | 15 | 21.9 | < 10 | 24 | 22.5 | 30 | 21 | 25 | 42
| 1 | 6 | 21.097 | 24.8 | 22 | 23.3 | 19.8 | 26.3 | 22 | 20.6 | 30 | 8 | 22 | 22.5 | 23 | 24 | 22.226 | 22 | 23.4 | 12 | 21.8 | < 10 | 22 | 23.1 | 25.3 | 18 | 24
| Mean |  | 23.0037 | 24.8833 | 22.1667 | 23.0833 | 20.3333 | 25.6167 | 21.8333 | 21.7667 | 28.3333 | 7.6667 | 22 | 22.9 | 22.6667 | 26.3333 | 24.7616 | 22.5667 | 23.7 | 11.6667 | 23.4833 |  | 24.1667 | 22.1833 | 25.8667 | 20.1667 | 24.6667
| Median |  | 23.4315 | 24.85 | 22 | 22.9 | 19.85 | 25.55 | 21.5 | 21.1 | 30 | 7.5 | 22 | 22.55 | 22.5 | 25.5 | 25.5748 | 22.7 | 23.55 | 12.5 | 23.2 |  | 24 | 22.3 | 25.5 | 19.5 | 24.5
| Std Dev. |  | 1.2339 | 0.31885 | 1.169 | 0.99079 | 1.3808 | 0.85186 | 1.169 | 1.5253 | 4.0825 | 1.633 | 2.7568 | 0.81731 | 1.2111 | 4.4121 | 2.1351 | 0.50859 | 0.53666 | 3.266 | 1.7657 |  | 1.7224 | 0.64317 | 2.2358 | 2.7869 | 0.8165
| Rel.Std.Dev. |  | 5.364% | 1.281% | 5.274% | 4.292% | 6.791% | 3.325% | 5.354% | 7.008% | 14.41% | 21.3% | 12.53% | 3.569% | 5.343% | 16.75% | 8.623% | 2.254% | 2.264% | 27.99% | 7.519% |  | 7.127% | 2.899% | 8.643% | 13.82% | 3.31%
| PDM3 |  | -0.1815% | 7.975% | -3.813% | 0.1642% | -11.77% | 11.16% | -5.26% | -5.549% | 22.95% | -66.73% | -4.537% | -0.6313% | -1.644% | 14.27% | 7.447% | -2.078% | 2.84% | -49.38% | 1.9% |  | 4.865% | -3.741% | 12.24% | -12.49% | 7.035%
| Z-Score (Absolute) |  | 0.06 | 1.06 | 0.39 | 0.1 | 1.37 | 1.45 | 0.57 | 0.6 | 2.89 | 8.11 | 0.48 | 0 | 0.12 | 1.83 | 0.99 | 0.18 | 0.43 | 5.98 | 0.31 | 6.87 | 0.67 | 0.38 | 1.58 | 1.46 | 0.94
| Analytical results for Pr in OREAS 501b (Certified Value 7.9 ppm) | [26 empty] | Certified
Pr | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | 2
3
| 1 | 1 | 01: 8.1751; 02: 7.998; 03: 7.1; 04: 8.262; 05: 8.3; 06: 29.6 | 1
| 1 | 2 | 01: 7.921; 02: 8.431; 03: 6.8; 04: 7.856; 05: 8.2; 06: 29.6 | #N/A
| 1 | 3 | 01: 8.4807; 02: 8.306; 03: 7.02; 04: 8.301; 05: 8.3; 06: 30.8 | 16
| 1 | 4 | 01: 8.6378; 02: 9.707; 03: 6.94; 04: 7.918; 05: 8; 06: 29.8 | 7.9034
| 1 | 5 | 01: 7.3283; 02: 7.988; 03: 7.04; 04: 8.023; 05: 8.4; 06: 31.2 | 43
| 1 | 6 | 01: 7.8687; 02: 8.465; 03: 6.64; 04: 8.064; 05: 8.1; 06: 31.4
| Mean |  | 01: 8.0686; 02: 8.4825; 03: 6.9233; 04: 8.0707; 05: 8.2167; 06: 30.4
| Median |  | 01: 8.048; 02: 8.3685; 03: 6.98; 04: 8.0435; 05: 8.25; 06: 30.3
| Std Dev. |  | 01: 0.47222; 02: 0.63447; 03: 0.17317; 04: 0.17965; 05: 0.1472; 06: 0.82946
| Rel.Std.Dev. |  | 01: 5.853%; 02: 7.48%; 03: 2.501%; 04: 2.226%; 05: 1.791%; 06: 2.728%
| PDM3 |  | 01: 2.091%; 02: 7.328%; 03: -12.4%; 04: 2.117%; 05: 3.964%; 06: 284.6%
| Z-Score (Absolute) |  | 01: 0.24; 02: 1.1; 03: 3.98; 04: 0.24; 05: 0.24; 06: 72.52
| Analytical results for Rb in OREAS 501b (Certified Value 184 ppm) | [26 empty] | Certified
Rb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [9 empty] | 0
0
| 1 | 1 | 211.85 | 197.24 | 183 | 199.5 | 178.5 | 172.5 | 202.44 | 162 | 183.1175 | 188 | 162.5 | 38 | 178 | 194.6 | 275 | 168.2 | [9 empty] | 1
| 1 | 2 | 210.52 | 198.56 | 170 | 195 | 171 | 169.5 | 203.55 | 161 | 178.7781 | 179.5 | 189.5 | 32 | 174 | 197.9 | 228 | 173.5 | [9 empty] | #N/A
| 1 | 3 | 205.3 | 199.58 | 184 | 198.5 | 180 | 166 | 208.16 | 171 | 180.2072 | 182 | 169 | 45 | 182.5 | 182.7 | 242 | 162.9 | [9 empty] | 16
| 1 | 4 | 205.02 | 207.95 | 192 | 203 | 185 | 162 | 211.53 | 164 | 183.1967 | 168 | 163 | 45 | 183 | 198.6 | 230 | 152 | [9 empty] | 184.2465
| 1 | 5 | 182.6 | 200.25 | 190 | 198.5 | 179.5 | 168.1 | 208.97 | 173.5 | 181.0918 | 179.5 | 161 | 43 | 181 | 199.1 | 275 | 176.7 | [9 empty] | 44
| 1 | 6 | 188.01 | 200.8 | 182 | 207 | 183.5 | 165.5 | 215.17 | 165.5 | 183.9306 | 177 | 178.5 | 43 | 163.5 | 197.7 | 256 | 153.8
| Mean |  | 200.55 | 200.73 | 183.5 | 200.25 | 179.5833 | 167.2667 | 208.3033 | 166.1667 | 181.7203 | 179 | 170.5833 | 41 | 177 | 195.1 | 251 | 164.5167
| Median |  | 205.16 | 199.915 | 183.5 | 199 | 179.75 | 167.05 | 208.565 | 164.75 | 182.1046 | 179.5 | 166 | 43 | 179.5 | 197.8 | 249 | 165.55
| Std Dev. |  | 12.24 | 3.757 | 7.7395 | 4.1803 | 4.893 | 3.6203 | 4.7958 | 5.0266 | 2.0178 | 6.5651 | 11.2934 | 5.099 | 7.4095 | 6.2747 | 21.1092 | 10.1669
| Rel.Std.Dev. |  | 6.103% | 1.872% | 4.218% | 2.088% | 2.725% | 2.164% | 2.302% | 3.025% | 1.11% | 3.668% | 6.62% | 12.44% | 4.186% | 3.216% | 8.41% | 6.18%
| PDM3 |  | 8.849% | 8.946% | -0.4051% | 8.686% | -2.531% | -9.216% | 13.06% | -9.813% | -1.371% | -2.848% | -7.416% | -77.75% | -3.933% | 5.891% | 36.23% | -10.71%
| Z-Score (Absolute) |  | 0.93 | 0.94 | 0.13 | 0.91 | 0.05 | 0.62 | 1.29 | 0.68 | 0.05 | 0.08 | 0.47 | 6.51 | 0.17 | 0.67 | 3.28 | 0.75
| Analytical results for Re in OREAS 501b (Indicative Value 0.004 ppm) | [26 empty] | Indicative
Re | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [13 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 17 | 18 | 20 | 23 | 25 | [13 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [13 empty] | 3
3
| 1 | 1 | 0.002 | < 0.1 | 0.004 | 0.004 | < 0.005 | < 0.002 | 0.003 | 0.003 | 0.008 | < 0.002 | < 0.01 | < 0.1 | [13 empty] | 1
| 1 | 2 | 0.003 | < 0.1 | 0.005 | 0.005 | 0.005 | 0.002 | 0.003 | 0.003 | 0.007 | 0.002 | < 0.01 | < 0.1 | [13 empty] | 4
| 1 | 3 | 0.004 | < 0.1 | 0.005 | 0.004 | < 0.005 | < 0.002 | 0.003 | 0.003 | 0.008 | 0.003 | < 0.01 | < 0.1 | [13 empty] | 16
| 1 | 4 | 0.005 | < 0.1 | 0.005 | 0.003 | 0.007 | 0.005 | 0.003 | 0.003 | 0.007 | 0.003 | < 0.01 | < 0.1 | [13 empty] | 0.0040926
| 1 | 5 | 0.004 | < 0.1 | 0.005 | 0.004 | < 0.005 | 0.003 | 0.002 | 0.003 | 0.007 | 0.002 | < 0.01 | 0.2 | [13 empty] | 10
| 1 | 6 | 0.003 | < 0.1 | 0.006 | 0.004 | 0.005 | 0.002 | 0.002 | 0.003 | 0.008 | < 0.002 | < 0.01 | < 0.1
| Mean |  | 0.0035 |  | 0.005 | 0.004 | 0.0056667 | 0.003 | 0.0026667 | 0.003 | 0.0075 | 0.0025 |  | 0.2
| Median |  | 0.0035 |  | 0.005 | 0.004 | 0.005 | 0.0025 | 0.003 | 0.003 | 0.0075 | 0.0025 |  | 0.2
| Std Dev. |  | 01: 0.0010488; 03: 0.00063246; 04: 0.00063246; 05: 0.0011547; 06: 0.0014142; 08: 0.0005164; 12: 4.7507e-19; 16: 0.00054772; 17: 0.00057735
| Rel.Std.Dev. |  | 01: 29.97%; 03: 12.65%; 04: 15.81%; 05: 20.38%; 06: 47.14%; 08: 19.36%; 12: 1.584e-14%; 16: 7.303%; 17: 23.09%
| PDM3 |  | -14.48% |  | 22.17% | -2.262% | 38.46% | -26.7% | -34.84% | -26.7% | 83.26% | -38.91% |  | 4787%
| Z-Score (Absolute) |  | 0.37 | 35.13 | 0.18 | 0.18 | 0.31 | 0.67 | 0.67 | 0.55 | 1.1 | 0.8 | 2.02 | 41.26
| Analytical results for S in OREAS 501b (Certified Value 0.354 wt.%) | [26 empty] | Certified
S | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 10 | 11 | 12 | 13 | 15 | 17 | 18 | 19 | 20 | 24 | 25 | 26 | 27 | [5 empty] | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | [5 empty] | 3
3
| 1 | 1 | 0.38226 | 0.378 | 0.35 | 0.35 | 0.38 | 0.3 | 0.3468 | 0.342 | 0.4 | 0.37 | 0.3427 | 0.35 | 0.36 | 0.37 | 0.69 | 0.36 | 0.39 | 0.32 | 0.39 | 0.29 | [5 empty] | 1
| 1 | 2 | 0.37896 | 0.35 | 0.355 | 0.35 | 0.37 | 0.3 | 0.338 | 0.339 | 0.4 | 0.36 | 0.3439 | 0.36 | 0.35 | 0.38 | 0.65 | 0.36 | 0.38 | 0.32 | 0.4 | 0.29 | [5 empty] | #N/A
| 1 | 3 | 0.38034 | 0.372 | 0.35 | 0.35 | 0.37 | 0.3 | 0.3625 | 0.334 | 0.4 | 0.37 | 0.3355 | 0.34 | 0.34 | 0.37 | 0.58 | 0.35 | 0.38 | 0.31 | 0.37 | 0.27 | [5 empty] | 16
| 1 | 4 | 0.38388 | 0.361 | 0.355 | 0.36 | 0.37 | 0.3 | 0.3546 | 0.338 | 0.3 | 0.37 | 0.3302 | 0.35 | 0.35 | 0.36 | 0.6 | 0.35 | 0.37 | 0.32 | 0.39 | 0.3 | [5 empty] | 0.35356
| 1 | 5 | 0.34576 | 0.386 | 0.36 | 0.35 | 0.37 | 0.3 | 0.3681 | 0.331 | 0.4 | 0.37 | 0.3304 | 0.35 | 0.36 | 0.36 | 0.73 | 0.36 | 0.39 | 0.31 | 0.37 | 0.29 | [5 empty] | 45
| 1 | 6 | 0.38364 | 0.343 | 0.35 | 0.37 | 0.37 | 0.3 | 0.357 | 0.34 | 0.4 | 0.36 | 0.3367 | 0.36 | 0.34 | 0.36 | 0.66 | 0.35 | 0.4 | 0.31 | 0.38 | 0.29
| Mean |  | 0.37581 | 0.365 | 0.35333 | 0.355 | 0.37167 | 0.3 | 0.3545 | 0.33733 | 0.38333 | 0.36667 | 0.33657 | 0.35167 | 0.35 | 0.36667 | 0.65167 | 0.355 | 0.385 | 0.315 | 0.38333 | 0.28833
| Median |  | 0.3813 | 0.3665 | 0.3525 | 0.35 | 0.37 | 0.3 | 0.3558 | 0.3385 | 0.4 | 0.37 | 0.3361 | 0.35 | 0.35 | 0.365 | 0.655 | 0.355 | 0.385 | 0.315 | 0.385 | 0.29
| Std Dev. |  | 0.014841 | 0.016637 | 0.0040825 | 0.0083666 | 0.0040825 | 0 | 0.010833 | 0.0040825 | 0.040825 | 0.005164 | 0.00585 | 0.0075277 | 0.0089443 | 0.008165 | 0.055648 | 0.0054772 | 0.010488 | 0.0054772 | 0.012111 | 0.0098319
| Rel.Std.Dev. |  | 3.949% | 4.558% | 1.155% | 2.357% | 1.098% | 0% | 3.056% | 1.21% | 10.65% | 1.408% | 1.738% | 2.141% | 2.556% | 2.227% | 8.539% | 1.543% | 2.724% | 1.739% | 3.159% | 3.41%
| PDM3 |  | 6.293% | 3.237% | -0.06273% | 0.4087% | 5.123% | -15.15% | 0.2672% | -4.588% | 8.423% | 3.708% | -4.805% | -0.5341% | -1.006% | 3.708% | 84.32% | 0.4087% | 8.894% | -10.9% | 8.423% | -18.45%
| Z-Score (Absolute) |  | 0.82 | 0.39 | 0.07 | 0 | 0.65 | 2.16 | 0.02 | 0.69 | 1.11 | 0.46 | 0.72 | 0.13 | 0.2 | 0.46 | 11.65 | 0 | 1.18 | 1.57 | 1.11 | 2.62
| Analytical results for Sb in OREAS 501b (Certified Value 0.87 ppm) | [26 empty] | Certified
Sb | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES | 2
2
| 1 | 1 | 0.971 | 0.83 | 0.8 | 1 | 0.7 | 0.9 | 6 | 0.84 | < 10 | < 5 | < 5 | 0.8 | 12 | < 5 | 1.075 | 0.85 | 0.93 | < 5 | 0.79 | < 10 | 1.1 | 0.94 | 0.98 | < 2 | < 5 | 1
| 1 | 2 | 0.9898 | 0.83 | 0.6 | 0.82 | 1.16 | 0.9 | < 5 | 0.84 | < 10 | < 5 | < 5 | 0.79 | 11 | < 5 | 0.977 | 0.79 | 1 | < 5 | 0.81 | < 10 | 1.2 | 0.92 | 0.96 | < 2 | < 5 | #N/A
| 1 | 3 | 0.933 | 0.82 | 0.6 | 0.84 | 0.74 | 1 | 7 | 0.87 | < 10 | < 5 | < 5 | 0.82 | 12 | < 5 | 1.017 | 0.78 | 0.92 | < 5 | 0.81 | < 10 | 1 | 0.9 | 1.2 | < 2 | < 5 | 16
| 1 | 4 | 0.9928 | 0.85 | 0.8 | 0.9 | 0.72 | 0.9 | 6 | 0.84 | < 10 | < 5 | < 5 | 0.79 | 10 | < 5 | 0.949 | 0.83 | 0.96 | < 5 | 0.85 | < 10 | 1.1 | 0.93 | 1.13 | < 2 | < 5 | 0.87379
| 1 | 5 | 1.0473 | 0.84 | 0.8 | 0.84 | 0.7 | 0.9 | < 5 | 0.84 | < 10 | < 5 | < 5 | 0.83 | 10 | < 5 | 0.939 | 0.85 | 0.92 | < 5 | 0.76 | < 10 | 1.1 | 0.93 | 2.06 | < 2 | < 5 | 46
| 1 | 6 | 0.9363 | 0.86 | 0.8 | 0.96 | 0.76 | 0.9 | < 5 | 0.85 | < 10 | < 5 | < 5 | 0.81 | 12 | < 5 | 0.923 | 0.78 | 0.9 | < 5 | 0.76 | < 10 | 1 | 1.1 | 3.86 | < 2 | < 5
| Mean |  | 0.97837 | 0.83833 | 0.73333 | 0.89333 | 0.79667 | 0.91667 | 6.3333 | 0.84667 |  |  |  | 0.80667 | 11.1667 |  | 0.98 | 0.81333 | 0.93833 |  | 0.79667 |  | 1.0833 | 0.95333 | 1.6983
| Median |  | 0.9804 | 0.835 | 0.8 | 0.87 | 0.73 | 0.9 | 6 | 0.84 |  |  |  | 0.805 | 11.5 |  | 0.963 | 0.81 | 0.925 |  | 0.8 |  | 1.1 | 0.93 | 1.165
| Std Dev. |  | 0.042362 | 0.01472 | 0.10328 | 0.073394 | 0.17952 | 0.040825 | 0.57735 | 0.012111 |  |  |  | 0.01633 | 0.98319 |  | 0.057086 | 0.033862 | 0.036009 |  | 0.034448 |  | 0.075277 | 0.073121 | 1.1346
| Rel.Std.Dev. |  | 4.33% | 1.756% | 14.08% | 8.216% | 22.53% | 4.454% | 9.116% | 1.43% |  |  |  | 2.024% | 8.805% |  | 5.825% | 4.163% | 3.838% |  | 4.324% |  | 6.949% | 7.67% | 66.8%
| PDM3 |  | 11.97% | -4.058% | -16.07% | 2.237% | -8.826% | 4.907% | 624.8% | -3.104% |  |  |  | -7.682% | 1178% |  | 12.16% | -6.919% | 7.387% |  | -8.826% |  | 23.98% | 9.103% | 94.36%
| Z-Score (Absolute) |  | 0.01 | 0.52 | 0.91 | 0.32 | 0.67 | 0.23 | 17.24 | 0.49 | 33.18 | 14.79 | 14.79 | 0.64 | 37.47 | 14.79 | 0 | 0.61 | 0.15 | 14.79 | 0.67 | 33.18 | 0.38 | 0.1 | 2.64 | 3.75 | 14.79
| Analytical results for Sc in OREAS 501b (Certified Value 12.9 ppm) | [26 empty] | Certified
Sc | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  | 1
2
| 1 | 1 | 13.4191 | 13 | 13 | 12.6 | 11.5 | 12 | 12.9 | 13 | 12.8 | 13 | 12 | 14 | 12 | 13 | 12.8668 | 14.4 | 13.1 | 11 | 12.8 | 14 | 13.4 | 17.4 | 11 | 10 |  | 1
| 1 | 2 | 13.1052 | 14 | 13 | 12.5 | 11.2 | 12 | 12.7 | 13 | 12.8 | 13 | 12 | 13 | 12 | 13 | 12.9245 | 14.2 | 14.7 | 10 | 13.1 | 14 | 12.9 | 16.6 | 12 | 9 |  | #N/A
| 1 | 3 | 13.8942 | 14 | 13 | 12.8 | 11.5 | 12 | 12.8 | 14 | 12.9 | 12 | 12 | 13 | 12 | 12 | 13.3198 | 14.4 | 13 | 10 | 13 | 13 | 13.5 | 16.2 | 11 | 9 |  | 16
| 1 | 4 | 13.8285 | 14 | 13 | 13.9 | 11.4 | 12 | 12.8 | 14 | 13.1 | 13 | 9 | 13 | 12 | 12 | 13.2603 | 13.8 | 13.4 | 10 | 13.4 | 14 | 13.7 | 15.5 | 11 | 10 |  | 12.8978
| 1 | 5 | 12.2434 | 13 | 13 | 13 | 11 | 12 | 12.5 | 14 | 13.1 | 13 | 12 | 13 | 12 | 13 | 12.7509 | 14.7 | 13 | 11 | 12.9 | 14 | 14.1 | 47.1 | 12 | 10 |  | 47
| 1 | 6 | 13.2752 | 14 | 13 | 13.5 | 11.7 | 12 | 12.6 | 14 | 13.2 | 13 | 12 | 13 | 12 | 13 | 12.7798 | 14 | 13.3 | 11 | 12.9 | 14 | 14.5 | 17.2 | 11 | 10
| Mean |  | 13.2943 | 13.6667 | 13 | 13.05 | 11.3833 | 12 | 12.7167 | 13.6667 | 12.9833 | 12.8333 | 11.5 | 13.1667 | 12 | 12.6667 | 12.9837 | 14.25 | 13.4167 | 10.5 | 13.0167 | 13.8333 | 13.6833 | 21.6667 | 11.3333 | 9.6667
| Median |  | 13.3471 | 14 | 13 | 12.9 | 11.45 | 12 | 12.75 | 14 | 13 | 13 | 12 | 13 | 12 | 13 | 12.8956 | 14.3 | 13.2 | 10.5 | 12.95 | 14 | 13.6 | 16.9 | 11 | 10
| Std Dev. |  | 0.60022 | 0.5164 | 0 | 0.54681 | 0.24833 | 0 | 0.1472 | 0.5164 | 0.17224 | 0.40825 | 1.2247 | 0.40825 | 0 | 0.5164 | 0.24591 | 0.32094 | 0.64936 | 0.54772 | 0.2137 | 0.40825 | 0.56006 | 12.4787 | 0.5164 | 0.5164
| Rel.Std.Dev. |  | 4.515% | 3.779% | 0% | 4.19% | 2.182% | 0% | 1.158% | 3.779% | 1.327% | 3.181% | 10.65% | 3.101% | 0% | 4.077% | 1.894% | 2.252% | 4.84% | 5.216% | 1.642% | 2.951% | 4.093% | 57.59% | 4.556% | 5.342%
| PDM3 |  | 3.074% | 5.961% | 0.7921% | 1.18% | -11.74% | -6.961% | -1.405% | 5.961% | 0.6629% | -0.5001% | -10.84% | 2.084% | -6.961% | -1.792% | 0.6655% | 10.48% | 4.023% | -18.59% | 0.9213% | 7.253% | 6.09% | 67.99% | -12.13% | -25.05%
| Z-Score (Absolute) |  | 0.3 | 0.67 | 0.01 | 0.06 | 1.61 | 0.99 | 0.27 | 0.67 | 0.01 | 0.16 | 1.49 | 0.17 | 0.99 | 0.32 | 0.01 | 1.26 | 0.42 | 2.49 | 0.02 | 0.84 | 0.69 | 8.67 | 1.66 | 3.32
| Analytical results for Se in OREAS 501b (Indicative Value 3.63 ppm) | [26 empty] | Indicative
Se | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [9 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 16 | 17 | 18 | 19 | 20 | 22 | 25 | 26 | [9 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | [9 empty] | 2
2
| 1 | 1 | 3.901 | 3 | < 5 | 4 | 4 | 3 | 2.4 | 4 | 2.9785 | 4 | 4 | < 5 | 4 | < 10 | 5 | < 12 | [9 empty] | 1
| 1 | 2 | 3.8815 | 3 | < 5 | 4 | 4 | 3 | 2.6 | 5 | 2.9972 | 4 | 4 | < 5 | 3 | < 10 | 4 | < 12 | [9 empty] | 5
| 1 | 3 | 3.8766 | 3 | < 5 | 4 | 4 | 2 | 2.6 | 5 | 2.824 | 4 | 4 | < 5 | 3 | < 10 | 5 | < 12 | [9 empty] | 16
| 1 | 4 | 3.6873 | 3 | < 5 | 4 | 4 | 2 | 2.6 | 5 | 2.9599 | 4 | 3 | < 5 | 3 | < 10 | 6 | < 12 | [9 empty] | 3.6252
| 1 | 5 | 3.768 | 2 | < 5 | 4 | 4 | 2 | 2.5 | 5 | 2.9148 | 4 | 4 | < 5 | 3 | < 10 | 6 | < 12 | [9 empty] | 11
| 1 | 6 | 3.7419 | 3 | < 5 | 4 | 4 | 2 | 2.8 | 4 | 2.9869 | 4 | 4 | < 5 | 3 | < 10 | 5 | < 12
| Mean |  | 01: 3.8094; 02: 2.8333; 04: 4; 05: 4; 06: 2.3333; 07: 2.5833; 08: 4.6667; 09: 2.9435; 10: 4; 11: 3.8333; 13: 3.1667; 16: 5.1667
| Median |  | 01: 3.8223; 02: 3; 04: 4; 05: 4; 06: 2; 07: 2.6; 08: 5; 09: 2.9692; 10: 4; 11: 4; 13: 3; 16: 5
| Std Dev. |  | 01: 0.088638; 02: 0.40825; 04: 0; 05: 0; 06: 0.5164; 07: 0.13292; 08: 0.5164; 09: 0.065354; 10: 0; 11: 0.40825; 13: 0.40825; 16: 0.75277
| Rel.Std.Dev. |  | 01: 2.327%; 02: 14.41%; 04: 0%; 05: 0%; 06: 22.13%; 07: 5.145%; 08: 11.07%; 09: 2.22%; 10: 0%; 11: 10.65%; 13: 12.89%; 16: 14.57%
| PDM3 |  | 01: 5.079%; 02: -21.84%; 04: 10.34%; 05: 10.34%; 06: -35.64%; 07: -28.74%; 08: 28.73%; 09: -18.8%; 10: 10.34%; 11: 5.74%; 13: -12.65%; 16: 42.52%
| Z-Score (Absolute) |  | 0.13 | 0.79 | 0.67 | 0 | 0 | 1.12 | 0.96 | 0.45 | 0.71 | 0 | 0.11 | 0.67 | 0.56 | 4.05 | 0.79 | 5.39
| Analytical results for Si in OREAS 501b (Indicative Value 29.57 wt.%) | [26 empty] | Indicative
Si | Round | Replicate | 01: Lab |  | 1
| No | No | 01: 03 |  | wt.%
|  |  | 01: 4A*MS |  | 2
2
| 1 | 1 | 01: 29.6 |  | 1
| 1 | 2 | 01: 29.5 |  | 3
| 1 | 3 | 01: 29.7 |  | 16
| 1 | 4 | 01: 29.3 |  | 29.5667
| 1 | 5 | 01: 29.8 |  | 9
| 1 | 6 | 01: 29.5
| Mean |  | 01: 29.5667
| Median |  | 01: 29.55
| Std Dev. |  | 01: 0.17512
| Rel.Std.Dev. |  | 01: 0.5923%
| PDM3 |  | 01: -1.11e-13%
Z-Score (Absolute) | bar: 01
| Analytical results for Sm in OREAS 501b (Certified Value 5.97 ppm) | [26 empty] | Certified
Sm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab |  | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 |  | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS |  | 2
3
| 1 | 1 | 01: 6.0531; 02: 6.19; 03: 5.25; 04: 5.85; 05: 6; 06: 5.6 |  | 1
| 1 | 2 | 01: 5.9355; 02: 6.42; 03: 5.1; 04: 5.7; 05: 6.1; 06: 5.6 |  | 9
| 1 | 3 | 01: 6.4636; 02: 6.4; 03: 5.15; 04: 6.01; 05: 5.7; 06: 5.7 |  | 16
| 1 | 4 | 01: 6.0045; 02: 6.61; 03: 5.25; 04: 5.68; 05: 6.2; 06: 5.6 |  | 5.9725
| 1 | 5 | 01: 5.3848; 02: 6.32; 03: 5.2; 04: 5.72; 05: 6.4; 06: 5.7 |  | 48
| 1 | 6 | 01: 5.6824; 02: 6.35; 03: 4.9; 04: 5.9; 05: 6.1; 06: 5.8
| Mean |  | 01: 5.9207; 02: 6.3817; 03: 5.1417; 04: 5.81; 05: 6.0833; 06: 5.6667
| Median |  | 01: 5.97; 02: 6.375; 03: 5.175; 04: 5.785; 05: 6.1; 06: 5.65
| Std Dev. |  | 01: 0.36428; 02: 0.13819; 03: 0.13197; 04: 0.13176; 05: 0.23166; 06: 0.08165
| Rel.Std.Dev. |  | 01: 6.153%; 02: 2.165%; 03: 2.567%; 04: 2.268%; 05: 3.808%; 06: 1.441%
| PDM3 |  | 01: -0.8675%; 02: 6.852%; 03: -13.91%; 04: -2.72%; 05: 1.856%; 06: -5.12%
| Z-Score (Absolute) |  | 01: 0.18; 02: 1.67; 03: 2.34; 04: 0.18; 05: 0.71; 06: 0.64
| Analytical results for Sn in OREAS 501b (Certified Value 5.58 ppm) | [26 empty] | Certified
Sn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | 27 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 2
3
| 1 | 1 | 6.3965 | 5.3 | 6 | 5.8 | 4.5 | 5.9 | < 10 | 5.7 | < 15 | < 10 | 5 | 5.1 | < 5 | 6.12 | 5.7 | 6 | < 10 | 5.4 | 5.2 | 7.8 | < 5 | < 10 |  |  |  | 1
| 1 | 2 | 6.3513 | 5.6 | 5 | 5.4 | 4.9 | 5.9 | < 10 | 5.8 | < 15 | < 10 | 6 | 5.2 | < 5 | 5.96 | 5.5 | 6.5 | < 10 | 5.2 | 5.2 | 7.1 | < 5 | < 10 |  |  |  | #N/A
| 1 | 3 | 6.0612 | 5.5 | 6 | 5.4 | 4.8 | 5.7 | < 10 | 5.8 | < 15 | < 10 | 6 | 5.2 | < 5 | 6.01 | 5.4 | 6.1 | < 10 | 5.3 | 4.9 | 7.2 | < 5 | < 10 |  |  |  | 16
| 1 | 4 | 6.2026 | 5.8 | 6 | 5.9 | 4.7 | 5.8 | < 10 | 5.8 | < 15 | < 10 | 5 | 5.1 | < 5 | 5.91 | 5.4 | 6.5 | < 10 | 5.4 | 5.4 | 7 | < 5 | < 10 |  |  |  | 5.5767
| 1 | 5 | 5.6333 | 5.3 | 6 | 5.5 | 4.8 | 6.1 | < 10 | 5.7 | < 15 | < 10 | 6 | 5.2 | < 5 | 5.83 | 5.5 | 6.1 | < 10 | 5.2 | 5.5 | 17.9 | < 5 | < 10 |  |  |  | 49
| 1 | 6 | 5.7143 | 5.5 | 5 | 5.8 | 5 | 5.8 | < 10 | 5.8 | < 15 | < 10 | 7 | 5.2 | < 5 | 5.61 | 5.4 | 6 | < 10 | 4.9 | 5.3 | 7.5 | < 5 | < 10
| Mean |  | 6.0599 | 5.5 | 5.6667 | 5.6333 | 4.7833 | 5.8667 |  | 5.7667 |  |  | 5.8333 | 5.1667 |  | 5.9067 | 5.4833 | 6.2 |  | 5.2333 | 5.25 | 9.0833
| Median |  | 6.1319 | 5.5 | 6 | 5.65 | 4.8 | 5.85 |  | 5.8 |  |  | 6 | 5.2 |  | 5.935 | 5.45 | 6.1 |  | 5.25 | 5.25 | 7.35
| Std Dev. |  | 0.32249 | 0.18974 | 0.5164 | 0.22509 | 0.17224 | 0.13663 |  | 0.05164 |  |  | 0.75277 | 0.05164 |  | 0.17489 | 0.1169 | 0.23664 |  | 0.18619 | 0.20736 | 4.3292
| Rel.Std.Dev. |  | 5.322% | 3.45% | 9.113% | 3.996% | 3.601% | 2.329% |  | 0.8955% |  |  | 12.9% | 0.9995% |  | 2.961% | 2.132% | 3.817% |  | 3.558% | 3.95% | 47.66%
| PDM3 |  | 8.665% | -1.375% | 1.614% | 1.016% | -14.23% | 5.2% |  | 3.407% |  |  | 4.603% | -7.352% |  | 5.918% | -1.673% | 11.18% |  | -6.156% | -5.858% | 62.88%
| Z-Score (Absolute) |  | 0.31 | 0.36 | 0.16 | 0.2 | 1.23 | 0.08 | 5.07 | 0.04 | 11.11 | 5.07 | 0.04 | 0.76 | 0.97 | 0.13 | 0.38 | 0.48 | 5.07 | 0.68 | 0.66 | 3.97 | 0.97 | 5.07
| Analytical results for Sr in OREAS 501b (Certified Value 338 ppm) | [26 empty] | Certified
Sr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 350.1 | 323.72 | 320 | 340 | 350 | 344 | 334 | 335.15 | 322 | 340 | 375 | 360 | 286 | 348 | 327.4143 | 359 | 350 | 206 | 345 | 267 | 336 | 310 | 347.2 | 229 | 260 | 1
| 1 | 2 | 331.834 | 327.21 | 299 | 336 | 350 | 341 | 334 | 347.09 | 327 | 339 | 364 | 343 | 295 | 350 | 327.6487 | 341 | 366 | 184 | 344 | 256 | 333 | 294 | 346.9 | 160 | 260 | #N/A
| 1 | 3 | 363.739 | 327.27 | 307 | 332 | 351 | 320 | 327 | 356.65 | 328 | 323 | 368 | 347 | 288 | 339 | 332.236 | 340 | 344 | 169 | 329 | 265 | 333 | 297 | 345.2 | 170 | 240 | 16
| 1 | 4 | 359.148 | 327.61 | 318 | 348 | 348 | 335 | 335 | 356.96 | 332 | 343 | 316 | 352 | 285 | 341 | 328.7182 | 339 | 346 | 171 | 327 | 263 | 332 | 288 | 334.9 | 161 | 280 | 337.6362
| 1 | 5 | 320.206 | 325.3 | 313 | 336 | 349 | 333 | 329 | 357.87 | 329 | 337 | 380 | 348 | 282 | 344 | 328.633 | 357 | 344 | 240 | 349 | 259 | 339 | 298 | 320.5 | 189 | 260 | 50
| 1 | 6 | 344.306 | 323.15 | 306 | 356 | 352 | 331 | 341 | 351.74 | 330 | 330 | 374 | 343 | 286 | 352 | 317.9191 | 335 | 338 | 203 | 337 | 266 | 331 | 311 | 320.5 | 160 | 260
| Mean |  | 344.8888 | 325.71 | 310.5 | 341.3333 | 350 | 334 | 333.3333 | 350.91 | 328 | 335.3333 | 362.8333 | 348.8333 | 287 | 345.6667 | 327.0949 | 345.1667 | 348 | 195.5 | 338.5 | 262.6667 | 334 | 299.6667 | 335.8667 | 178.1667 | 260
| Median |  | 347.203 | 326.255 | 310 | 338 | 350 | 334 | 334 | 354.195 | 328.5 | 338 | 371 | 347.5 | 286 | 346 | 328.1408 | 340.5 | 345 | 193.5 | 340.5 | 264 | 333 | 297.5 | 340.05 | 165.5 | 260
| Std Dev. |  | 16.5277 | 1.9481 | 7.9687 | 9.0037 | 1.4142 | 8.438 | 4.9261 | 8.7367 | 3.4059 | 7.4476 | 23.6171 | 6.4317 | 4.3818 | 5.164 | 4.8174 | 10.1669 | 9.6333 | 26.7638 | 9.0277 | 4.3205 | 2.9665 | 9.0921 | 12.7318 | 27.2867 | 12.6491
| Rel.Std.Dev. |  | 4.792% | 0.5981% | 2.566% | 2.638% | 0.4041% | 2.526% | 1.478% | 2.49% | 1.038% | 2.221% | 6.509% | 1.844% | 1.527% | 1.494% | 1.473% | 2.946% | 2.768% | 13.69% | 2.667% | 1.645% | 0.8882% | 3.034% | 3.791% | 15.32% | 4.865%
| PDM3 |  | 2.148% | -3.532% | -8.037% | 1.095% | 3.662% | -1.077% | -1.274% | 3.931% | -2.854% | -0.6821% | 7.463% | 3.316% | -15% | 2.378% | -3.122% | 2.23% | 3.07% | -42.1% | 0.2558% | -22.2% | -1.077% | -11.25% | -0.5241% | -47.23% | -22.99%
| Z-Score (Absolute) |  | 0.63 | 0.48 | 1.36 | 0.42 | 0.92 | 0 | 0.04 | 0.98 | 0.35 | 0.08 | 1.67 | 0.86 | 2.72 | 0.67 | 0.4 | 0.65 | 0.81 | 8.01 | 0.26 | 4.12 | 0 | 1.98 | 0.11 | 9.01 | 4.28
| Analytical results for Ta in OREAS 501b (Certified Value 1.38 ppm) | [26 empty] | Certified
Ta | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 10 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 27 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | [7 empty] | 2
2
| 1 | 1 | 1.1251 | 1.33 | 1 | 1.33 | 1.05 | 1.2 | 1.54 | < 5 | 1.41 | 2.0235 | 1.57 | 1.42 | < 1 | 1.4 | 1.3 | 1.83 | 1.33 | < 5 | [7 empty] | 1
| 1 | 2 | 1.0681 | 1.35 | 0.9 | 1.28 | 1.07 | 1.3 | 1.54 | < 5 | 1.42 | 1.8149 | 1.54 | 1.49 | 1 | 1.39 | 1.4 | 1.71 | 0.91 | < 5 | [7 empty] | 10
| 1 | 3 | 1.0662 | 1.35 | 0.9 | 1.33 | 1.04 | 1.2 | 1.5 | < 5 | 1.47 | 1.7799 | 1.52 | 1.38 | < 1 | 1.44 | 1.3 | 1.64 | 1.92 | < 5 | [7 empty] | 16
| 1 | 4 | 1.1071 | 1.38 | 1 | 1.36 | 1.02 | 1.3 | 1.5 | < 5 | 1.51 | 1.6957 | 1.52 | 1.45 | < 1 | 1.37 | 1.5 | 1.78 | 1.53 | < 5 | [7 empty] | 1.3833
| 1 | 5 | 1.0354 | 1.33 | 0.9 | 1.3 | 1.02 | 1.3 | 1.48 | < 5 | 1.41 | 1.6732 | 1.58 | 1.39 | < 1 | 1.3 | 1.7 | 1.9 | 3.17 | < 5 | [7 empty] | 51
| 1 | 6 | 0.9863 | 1.3 | 1 | 1.4 | 1.07 | 1.3 | 1.46 | < 5 | 1.46 | 1.6048 | 1.53 | 1.4 | < 1 | 1.28 | 1.3 | 1.96 | 1.49 | < 5
| Mean |  | 1.0647 | 1.34 | 0.95 | 1.3333 | 1.045 | 1.2667 | 1.5033 |  | 1.4467 | 1.7653 | 1.5433 | 1.4217 | 1 | 1.3633 | 1.4167 | 1.8033 | 1.725
| Median |  | 1.0672 | 1.34 | 0.95 | 1.33 | 1.045 | 1.3 | 1.5 |  | 1.44 | 1.7378 | 1.535 | 1.41 | 1 | 1.38 | 1.35 | 1.805 | 1.51
| Std Dev. |  | 0.04993 | 0.026833 | 0.054772 | 0.04274 | 0.022583 | 0.05164 | 0.032042 |  | 0.040332 | 0.14725 | 0.02582 | 0.041673 |  | 0.061536 | 0.16021 | 0.11877 | 0.77969
| Rel.Std.Dev. |  | 4.69% | 2.002% | 5.766% | 3.205% | 2.161% | 4.077% | 2.131% |  | 2.788% | 8.341% | 1.673% | 2.931% |  | 4.514% | 11.31% | 6.586% | 45.2%
| PDM3 |  | -23.03% | -3.13% | -31.32% | -3.612% | -24.46% | -8.431% | 8.678% |  | 4.581% | 27.62% | 11.57% | 2.774% | -27.71% | -1.443% | 2.413% | 30.37% | 24.7%
| Z-Score (Absolute) |  | 1.04 | 0.23 | 1.38 | 0.25 | 1.1 | 0.45 | 0.25 | 10.54 | 0.08 | 1.02 | 0.37 | 0.01 | 1.23 | 0.16 | 0.01 | 1.13 | 0.9 | 10.54
| Analytical results for Tb in OREAS 501b (Certified Value 0.79 ppm) | [26 empty] | Certified
Tb | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 16; 06: 23; 07: 24; 08: 25 | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*OES/MS; 06: 4A*MS; 07: 4A*MS; 08: 4A*OES/MS | 2
3
| 1 | 1 | 01: 0.827; 02: 0.808; 03: 0.7; 04: 0.802; 05: 0.84623; 06: 0.8; 07: 0.86; 08: 0.7 | 1
| 1 | 2 | 01: 0.8146; 02: 0.805; 03: 0.68; 04: 0.766; 05: 0.84042; 06: 0.8; 07: 0.87; 08: 0.7 | #N/A
| 1 | 3 | 01: 0.8634; 02: 0.808; 03: 0.7; 04: 0.778; 05: 0.8468; 06: 0.8; 07: 0.87; 08: 0.7 | 16
| 1 | 4 | 01: 0.8572; 02: 0.873; 03: 0.7; 04: 0.763; 05: 0.85127; 06: 0.8; 07: 0.85; 08: 0.67 | 0.78936
| 1 | 5 | 01: 0.757; 02: 0.804; 03: 0.72; 04: 0.755; 05: 0.83895; 06: 0.8; 07: 0.89; 08: 0.62 | 52
| 1 | 6 | 01: 0.7871; 02: 0.809; 03: 0.68; 04: 0.792; 05: 0.85444; 06: 0.8; 07: 0.9; 08: 0.72
| Mean |  | 01: 0.81772; 02: 0.81783; 03: 0.69667; 04: 0.776; 05: 0.84635; 06: 0.8; 07: 0.87333; 08: 0.685
| Median |  | 01: 0.8208; 02: 0.808; 03: 0.7; 04: 0.772; 05: 0.84652; 06: 0.8; 07: 0.87; 08: 0.7
| Std Dev. |  | 01: 0.040904; 02: 0.027096; 03: 0.015055; 04: 0.018144; 05: 0.0059976; 06: 1.2162e-16; 07: 0.018619; 08: 0.035637
| Rel.Std.Dev. |  | 01: 5.002%; 02: 3.313%; 03: 2.161%; 04: 2.338%; 05: 0.7086%; 06: 1.52e-14%; 07: 2.132%; 08: 5.202%
| PDM3 |  | 01: 3.593%; 02: 3.607%; 03: -11.74%; 04: -1.692%; 05: 7.22%; 06: 1.348%; 07: 10.64%; 08: -13.22%
| Z-Score (Absolute) |  | 01: 0.17; 02: 0.17; 03: 2.15; 04: 0.63; 05: 0.72; 06: 0.17; 07: 1.24; 08: 2.37
| Analytical results for Te in OREAS 501b (Certified Value 0.093 ppm) | [26 empty] | Certified
Te | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 10 | 12 | 13 | 16 | 17 | 18 | 19 | 20 | 23 | 25 | 26 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | [7 empty] | 3
3
| 1 | 1 | 0.3722 | 0.1 | < 0.2 | 0.1 | 0.08 | < 0.5 | 0.1 | < 10 | < 0.05 | < 5 | 0.14887 | 0.08 | 0.1 | < 1 | < 0.05 | < 1 | 0.1 | < 1 | [7 empty] | 1
| 1 | 2 | 0.5015 | 0.1 | 0.2 | 0.09 | 0.07 | < 0.5 | 0.1 | < 10 | < 0.05 | < 5 | 0.11331 | 0.08 | 0.11 | < 1 | 0.09 | < 1 | 0.14 | < 1 | [7 empty] | #N/A
| 1 | 3 | 0.4395 | < 0.1 | < 0.2 | 0.11 | 0.08 | < 0.5 | 0.09 | < 10 | < 0.05 | < 5 | 0.10312 | 0.08 | 0.07 | 4 | 0.08 | < 1 | < 0.05 | < 1 | [7 empty] | 16
| 1 | 4 | 0.4421 | 0.1 | < 0.2 | 0.08 | 0.06 | < 0.5 | 0.11 | < 10 | < 0.05 | < 5 | 0.09589 | 0.08 | 0.09 | 2 | 0.09 | < 1 | 0.09 | < 1 | [7 empty] | 0.093285
| 1 | 5 | 0.2221 | 0.1 | < 0.2 | 0.07 | 0.09 | < 0.5 | 0.11 | < 10 | < 0.05 | < 5 | 0.09176 | 0.09 | 0.07 | 3 | 0.06 | < 1 | 0.5 | < 1 | [7 empty] | 53
| 1 | 6 | 0.455 | 0.1 | < 0.2 | 0.09 | 0.11 | < 0.5 | 0.1 | < 10 | < 0.05 | < 5 | 0.090384 | 0.09 | 0.1 | 1 | 0.1 | < 1 | 0.18 | < 1
| Mean |  | 01: 0.4054; 02: 0.1; 03: 0.2; 04: 0.09; 05: 0.081667; 07: 0.10167; 11: 0.10722; 12: 0.083333; 13: 0.09; 15: 2.5; 16: 0.084; 18: 0.202
| Median |  | 01: 0.4408; 02: 0.1; 03: 0.2; 04: 0.09; 05: 0.08; 07: 0.1; 11: 0.099507; 12: 0.08; 13: 0.095; 15: 2.5; 16: 0.09; 18: 0.14
| Std Dev. |  | 01: 0.0989; 02: 0; 04: 0.014142; 05: 0.017224; 07: 0.0075277; 11: 0.02209; 12: 0.005164; 13: 0.016733; 15: 1.291; 16: 0.015166; 18: 0.17035
| Rel.Std.Dev. |  | 01: 24.4%; 02: 0%; 04: 15.71%; 05: 21.09%; 07: 7.404%; 11: 20.6%; 12: 6.197%; 13: 18.59%; 15: 51.64%; 16: 18.05%; 18: 84.33%
| PDM3 |  | 01: 334.6%; 02: 7.199%; 03: 114.4%; 04: -3.521%; 05: -12.45%; 07: 8.986%; 11: 14.94%; 12: -10.67%; 13: -3.521%; 15: 2580%; 16: -9.953%; 18: 116.5%
| Z-Score (Absolute) |  | 2.87 | 0.46 | 0.63 | 0.57 | 0.66 | 3.9 | 0.44 | 107.31 | 1 | 52.88 | 0.38 | 0.64 | 0.57 | 20.23 | 0.69 | 9.34 | 0.38 | 9.34
| Analytical results for Th in OREAS 501b (Certified Value 18 ppm) | [26 empty] | Certified
Th | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [7 empty] | 1
2
| 1 | 1 | 18.792 | 17.69 | 19.4 | 18 | 14.1 | 18 | 18.64 | 24 | 17.1 | 20 | 18.6925 | 19 | 16.6 | < 50 | 18.8 | 19.21 | 19.4 | 20.3 | [7 empty] | 1
| 1 | 2 | 20.053 | 18.19 | 18.9 | 16.7 | 14.9 | 18.7 | 16.5 | 22 | 16.4 | 20 | 17.4891 | 17.8 | 15.7 | < 50 | 18.5 | 17.6 | 19.7 | 18.8 | [7 empty] | #N/A
| 1 | 3 | 18.372 | 18.3 | 18.9 | 18.3 | 15.3 | 18.3 | 17.92 | 20 | 17.4 | 20 | 16.3471 | 18.6 | 15.2 | < 50 | 21.9 | 18.19 | 18.7 | 17 | [7 empty] | 16
| 1 | 4 | 18.415 | 18.67 | 18.5 | 19 | 15.4 | 17.5 | 16.75 | 15 | 18.4 | 20 | 17.8885 | 16 | 14.8 | < 50 | 16.5 | 17.94 | 20 | 17.6 | [7 empty] | 17.9915
| 1 | 5 | 17.397 | 17.63 | 18.2 | 17.8 | 14.5 | 20.3 | 17.19 | 22 | 20.4 | 20 | 16.8736 | 17.2 | 14.6 | < 50 | 18.4 | 17.64 | 21.5 | 13 | [7 empty] | 54
| 1 | 6 | 16.394 | 18.15 | 17.6 | 19 | 15.5 | 18.3 | 17.85 | 23 | 15.3 | 20 | 18.2848 | 18.9 | 17.1 | < 50 | 17.1 | 18.33 | 20.5 | 17.2
| Mean |  | 18.2372 | 18.105 | 18.5833 | 18.1333 | 14.95 | 18.5167 | 17.475 | 21 | 17.5 | 20 | 17.5959 | 17.9167 | 15.6667 |  | 18.5333 | 18.1517 | 19.9667 | 17.3167
| Median |  | 18.3935 | 18.17 | 18.7 | 18.15 | 15.1 | 18.3 | 17.52 | 22 | 17.25 | 20 | 17.6888 | 18.2 | 15.45 |  | 18.45 | 18.065 | 19.85 | 17.4
| Std Dev. |  | 1.2456 | 0.39099 | 0.63061 | 0.86178 | 0.55767 | 0.96003 | 0.80664 | 3.2249 | 1.7573 | 0 | 0.87811 | 1.1669 | 1.0033 |  | 1.8769 | 0.59402 | 0.96264 | 2.4466
| Rel.Std.Dev. |  | 6.83% | 2.16% | 3.393% | 4.752% | 3.73% | 5.185% | 4.616% | 15.36% | 10.04% | 0% | 4.99% | 6.513% | 6.404% |  | 10.13% | 3.273% | 4.821% | 14.13%
| PDM3 |  | 1.365% | 0.6308% | 3.289% | 0.7883% | -16.91% | 2.919% | -2.871% | 16.72% | -2.732% | 11.16% | -2.199% | -0.416% | -12.92% |  | 3.012% | 0.8902% | 10.98% | -3.751%
| Z-Score (Absolute) |  | 0.11 | 0.04 | 0.5 | 0.01 | 3.62 | 0.42 | 0.76 | 3.24 | 0.73 | 2.11 | 0.62 | 0.26 | 2.81 | 36.13 | 0.44 | 0.01 | 2.07 | 0.94
| Analytical results for Ti in OREAS 501b (Certified Value 0.466 wt.%) | [26 empty] | Certified
Ti | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | wt.%
|  |  | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 3
3
| 1 | 1 | 0.46568 | 0.4704 | 0.48 | 0.453 | 0.468 | 0.486 | 0.44 | 0.4673 | 0.43 | 0.48 | 0.52 | 0.482 | 0.4538 | 0.47 | 0.448 | 0.468 | 0.469 | 0.3168 | 0.47 | 0.3876 | 0.5 | 0.47 | 0.48 | 0.47 | 0.35 | 1
| 1 | 2 | 0.45717 | 0.4825 | 0.47 | 0.443 | 0.469 | 0.487 | 0.43 | 0.4671 | 0.44 | 0.49 | 0.51 | 0.466 | 0.4662 | 0.47 | 0.459 | 0.445 | 0.494 | 0.3 | 0.468 | 0.3756 | 0.5 | 0.45 | 0.49 | 0.48 | 0.36 | #N/A
| 1 | 3 | 0.46843 | 0.48 | 0.48 | 0.454 | 0.467 | 0.481 | 0.43 | 0.4763 | 0.44 | 0.48 | 0.51 | 0.465 | 0.4578 | 0.46 | 0.47 | 0.446 | 0.467 | 0.2977 | 0.465 | 0.384 | 0.52 | 0.45 | 0.51 | 0.48 | 0.33 | 16
| 1 | 4 | 0.46328 | 0.481 | 0.47 | 0.472 | 0.468 | 0.478 | 0.44 | 0.4769 | 0.44 | 0.48 | 0.46 | 0.464 | 0.4549 | 0.46 | 0.473 | 0.459 | 0.465 | 0.302 | 0.452 | 0.3853 | 0.5 | 0.45 | 0.45 | 0.49 | 0.38 | 0.46633
| 1 | 5 | 0.42979 | 0.4752 | 0.48 | 0.449 | 0.467 | 0.487 | 0.43 | 0.4774 | 0.45 | 0.49 | 0.52 | 0.468 | 0.4535 | 0.46 | 0.458 | 0.469 | 0.459 | 0.3303 | 0.473 | 0.38 | 0.51 | 0.45 | 0.44 | 0.48 | 0.36 | 55
| 1 | 6 | 0.44714 | 0.471 | 0.47 | 0.467 | 0.475 | 0.481 | 0.45 | 0.485 | 0.45 | 0.49 | 0.51 | 0.479 | 0.4594 | 0.48 | 0.453 | 0.442 | 0.454 | 0.3188 | 0.459 | 0.3905 | 0.49 | 0.47 | 0.43 | 0.47 | 0.36
| Mean |  | 0.45525 | 0.47668 | 0.475 | 0.45633 | 0.469 | 0.48333 | 0.43667 | 0.475 | 0.44167 | 0.485 | 0.505 | 0.47067 | 0.4576 | 0.46667 | 0.46017 | 0.45483 | 0.468 | 0.31093 | 0.4645 | 0.38383 | 0.50333 | 0.45667 | 0.46667 | 0.47833 | 0.35667
| Median |  | 0.46023 | 0.4776 | 0.475 | 0.4535 | 0.468 | 0.4835 | 0.435 | 0.4766 | 0.44 | 0.485 | 0.51 | 0.467 | 0.45635 | 0.465 | 0.4585 | 0.4525 | 0.466 | 0.3094 | 0.4665 | 0.38465 | 0.5 | 0.45 | 0.465 | 0.48 | 0.36
| Std Dev. |  | 0.014592 | 0.0052431 | 0.0054772 | 0.011021 | 0.0030332 | 0.0038297 | 0.008165 | 0.0068229 | 0.0075277 | 0.0054772 | 0.022583 | 0.0077889 | 0.0048121 | 0.008165 | 0.0096626 | 0.01209 | 0.013885 | 0.013006 | 0.0077653 | 0.0053534 | 0.010328 | 0.010328 | 0.031411 | 0.0075277 | 0.01633
| Rel.Std.Dev. |  | 3.205% | 1.1% | 1.153% | 2.415% | 0.6467% | 0.7924% | 1.87% | 1.436% | 1.704% | 1.129% | 4.472% | 1.655% | 1.052% | 1.75% | 2.1% | 2.658% | 2.967% | 4.183% | 1.672% | 1.395% | 2.052% | 2.262% | 6.731% | 1.574% | 4.578%
| PDM3 |  | -2.375% | 2.221% | 1.86% | -2.143% | 0.5736% | 3.647% | -6.36% | 1.86% | -5.288% | 4.005% | 8.293% | 0.931% | -1.871% | 0.0732% | -1.321% | -2.464% | 0.3591% | -33.32% | -0.3914% | -17.69% | 7.936% | -2.071% | 0.0732% | 2.575% | -23.52%
| Z-Score (Absolute) |  | 0.75 | 0.65 | 0.54 | 0.67 | 0.15 | 1.09 | 1.96 | 0.54 | 1.63 | 1.2 | 2.5 | 0.26 | 0.59 | 0 | 0.42 | 0.77 | 0.09 | 10.16 | 0.14 | 5.41 | 2.39 | 0.65 | 0 | 0.76 | 7.18
| Analytical results for Tl in OREAS 501b (Certified Value 0.9 ppm) | [26 empty] | Certified
Tl | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [6 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | [6 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*OES/MS | [6 empty] | 2
3
| 1 | 1 | 0.8886 | 0.97 | 0.9 | 0.94 | 0.79 | 0.9 | 0.97 | 0.84 | < 5 | < 10 | 0.99412 | 0.93 | 0.97 | < 2 | 0.89 | < 20 | 0.93 | 0.77 | 1.25 | [6 empty] | 1
| 1 | 2 | 0.832 | 0.98 | 0.8 | 0.92 | 0.82 | 0.9 | 0.95 | 0.88 | < 5 | < 10 | 0.99528 | 0.93 | 0.98 | < 2 | 0.84 | < 20 | 0.97 | 0.79 | 1.21 | [6 empty] | #N/A
| 1 | 3 | 0.8288 | 0.98 | 0.8 | 0.91 | 0.83 | 1 | 1 | 0.88 | < 5 | < 10 | 0.97967 | 0.9 | 0.94 | < 2 | 0.85 | < 20 | 0.91 | 0.77 | 1.25 | [6 empty] | 16
| 1 | 4 | 0.853 | 0.99 | 0.9 | 0.97 | 0.85 | 0.9 | 0.97 | 0.9 | < 5 | < 10 | 0.98924 | 0.91 | 0.98 | < 2 | 0.82 | < 20 | 0.99 | 0.74 | 1.19 | [6 empty] | 0.90091
| 1 | 5 | 0.9577 | 0.99 | 0.8 | 0.93 | 0.78 | 0.9 | 0.95 | 0.84 | < 5 | < 10 | 0.98783 | 0.95 | 0.99 | < 2 | 0.81 | < 20 | 0.96 | 0.76 | 2.28 | [6 empty] | 56
| 1 | 6 | 0.8423 | 0.98 | 0.8 | 0.96 | 0.83 | 0.9 | 1 | 0.92 | < 5 | < 10 | 0.98686 | 0.9 | 0.97 | < 2 | 0.83 | < 20 | 0.96 | 0.79 | 1.16
| Mean |  | 0.86707 | 0.98167 | 0.83333 | 0.93833 | 0.81667 | 0.91667 | 0.97333 | 0.87667 |  |  | 0.98883 | 0.92 | 0.97167 |  | 0.84 |  | 0.95333 | 0.77 | 1.39
| Median |  | 0.84765 | 0.98 | 0.8 | 0.935 | 0.825 | 0.9 | 0.97 | 0.88 |  |  | 0.98853 | 0.92 | 0.975 |  | 0.835 |  | 0.96 | 0.77 | 1.23
| Std Dev. |  | 0.049363 | 0.0075277 | 0.05164 | 0.023166 | 0.026583 | 0.040825 | 0.022509 | 0.032042 |  |  | 0.0056302 | 0.02 | 0.017224 |  | 0.028284 |  | 0.028752 | 0.018974 | 0.4374
| Rel.Std.Dev. |  | 5.693% | 0.7668% | 6.197% | 2.469% | 3.255% | 4.454% | 2.313% | 3.655% |  |  | 0.5694% | 2.174% | 1.773% |  | 3.367% |  | 3.016% | 2.464% | 31.47%
| PDM3 |  | -3.757% | 8.964% | -7.501% | 4.154% | -9.351% | 1.749% | 8.039% | -2.691% |  |  | 9.759% | 2.119% | 7.854% |  | -6.761% |  | 5.819% | -14.53% | 54.29%
| Z-Score (Absolute) |  | 0.67 | 0.22 | 0.94 | 0.12 | 1.07 | 0.29 | 0.16 | 0.6 | 31.63 | 70.71 | 0.28 | 0.26 | 0.14 | 8.18 | 0.89 | 148.88 | 0 | 1.43 | 3.41
| Analytical results for Tm in OREAS 501b (Certified Value 0.36 ppm) | [26 empty] | Certified
Tm | Round | Replicate | 01: Lab; 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab | [6 empty] | 1
| No | No | 01: 01; 02: 02; 03: 03; 04: 08; 05: 23; 06: 24 | [6 empty] | ppm
|  |  | 01: 4A*MS; 02: 4A*OES/MS; 03: 4A*MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*MS | [6 empty] | 2
3
| 1 | 1 | 01: 0.346; 02: 0.37; 03: 0.34; 04: 0.38; 05: 0.4; 06: 0.34 | [6 empty] | 1
| 1 | 2 | 01: 0.3284; 02: 0.38; 03: 0.32; 04: 0.38; 05: 0.4; 06: 0.34 | [6 empty] | 11
| 1 | 3 | 01: 0.3601; 02: 0.37; 03: 0.34; 04: 0.4; 05: 0.4; 06: 0.34 | [6 empty] | 16
| 1 | 4 | 01: 0.352; 02: 0.4; 03: 0.32; 04: 0.39; 05: 0.4; 06: 0.34 | [6 empty] | 0.36278
| 1 | 5 | 01: 0.3121; 02: 0.35; 03: 0.32; 04: 0.38; 05: 0.4; 06: 0.35 | [6 empty] | 57
| 1 | 6 | 01: 0.3215; 02: 0.41; 03: 0.34; 04: 0.4; 05: 0.4; 06: 0.34
| Mean |  | 01: 0.33668; 02: 0.38; 03: 0.33; 04: 0.38833; 05: 0.4; 06: 0.34167
| Median |  | 01: 0.3372; 02: 0.375; 03: 0.33; 04: 0.385; 05: 0.4; 06: 0.34
| Std Dev. |  | 01: 0.018832; 02: 0.021909; 03: 0.010954; 04: 0.0098319; 05: 6.0809e-17; 06: 0.0040825
| Rel.Std.Dev. |  | 01: 5.593%; 02: 5.766%; 03: 3.32%; 04: 2.532%; 05: 1.52e-14%; 06: 1.195%
| PDM3 |  | 01: -7.194%; 02: 4.747%; 03: -9.036%; 04: 7.044%; 05: 10.26%; 06: -5.82%
| Z-Score (Absolute) |  | 01: 0.63; 02: 0.5; 03: 0.81; 04: 0.72; 05: 1.02; 06: 0.5
| Analytical results for U in OREAS 501b (Certified Value 4.86 ppm) | [26 empty] | Certified
U | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [7 empty] | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 08 | 11 | 12 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | [7 empty] | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | [7 empty] | 2
3
| 1 | 1 | 6.1377 | 4.78 | 5 | 5 | 3.8 | 5.4 | 4.74 | < 20 | 4.5 | 10 | 5.6934 | 5 | 4.5 | 53 | 4.7 | 4.83 | 5.21 | 4.7 | [7 empty] | 1
| 1 | 2 | 6.5386 | 4.94 | 4.5 | 4.9 | 4.3 | 5.4 | 5.14 | < 20 | 4.1 | < 10 | 5.289 | 4.9 | 4.7 | 43 | 4.6 | 4.91 | 5.79 | 4.5 | [7 empty] | #N/A
| 1 | 3 | 5.7002 | 4.83 | 4.5 | 4.8 | 4 | 5.3 | 5.36 | < 20 | 4.6 | 10 | 4.8161 | 5.1 | 5.3 | 41 | 5.5 | 4.85 | 4.94 | 4.3 | [7 empty] | 16
| 1 | 4 | 5.537 | 5.32 | 4.5 | 4.9 | 4 | 5.2 | 4.64 | < 20 | 4.7 | < 10 | 4.7847 | 4.3 | 4.5 | 46 | 4.4 | 4.89 | 5.8 | 4.6 | [7 empty] | 4.8625
| 1 | 5 | 5.1464 | 4.89 | 5.1 | 4.6 | 4 | 6.7 | 4.88 | < 20 | 4.8 | < 10 | 4.7547 | 4.7 | 4.9 | 53 | 4.6 | 4.7 | 5.52 | 4.4 | [7 empty] | 58
| 1 | 6 | 5.0156 | 5.55 | 4.7 | 5 | 4 | 5.4 | 4.98 | < 20 | 4.4 | < 10 | 5.4115 | 5.2 | 4.8 | 39 | 4 | 4.86 | 5.42 | 4.5
| Mean |  | 5.6792 | 5.0517 | 4.7167 | 4.8667 | 4.0167 | 5.5667 | 4.9567 |  | 4.5167 | 10 | 5.1249 | 4.8667 | 4.7833 | 45.8333 | 4.6333 | 4.84 | 5.4467 | 4.5
| Median |  | 5.6186 | 4.915 | 4.6 | 4.9 | 4 | 5.4 | 4.93 |  | 4.55 | 10 | 5.0525 | 4.95 | 4.75 | 44.5 | 4.6 | 4.855 | 5.47 | 4.5
| Std Dev. |  | 0.58201 | 0.31045 | 0.27142 | 0.15055 | 0.16021 | 0.56095 | 0.26455 |  | 0.24833 | 0 | 0.39507 | 0.3266 | 0.29944 | 6.0139 | 0.49261 | 0.074297 | 0.33512 | 0.14142
| Rel.Std.Dev. |  | 10.25% | 6.145% | 5.754% | 3.094% | 3.989% | 10.08% | 5.337% |  | 5.498% | 0% | 7.709% | 6.711% | 6.26% | 13.12% | 10.63% | 1.535% | 6.153% | 3.143%
| PDM3 |  | 16.8% | 3.891% | -2.999% | 0.08601% | -17.39% | 14.48% | 1.937% |  | -7.112% | 105.7% | 5.396% | 0.08601% | -1.628% | 842.6% | -4.713% | -0.4624% | 12.01% | -7.455%
| Z-Score (Absolute) |  | 1.54 | 0.28 | 0.39 | 0.09 | 1.79 | 1.31 | 0.09 | 30.22 | 0.79 | 10.19 | 0.43 | 0.09 | 0.26 | 81.96 | 0.56 | 0.14 | 1.07 | 0.82
| Analytical results for V in OREAS 501b (Certified Value 121 ppm) | [26 empty] | Certified
V | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
0
| 1 | 1 | 122.55 | 119 | 120 | 122 | 120 | 123 | 113 | 124 | 125 | 119 | 142 | 125 | 117 | 119 | 115.1308 | 124 | 127 | 101 | 123 | 106 | 127 | 121 | 123 | 114 | 108 | 1
| 1 | 2 | 116.619 | 124 | 125 | 120 | 120 | 123 | 110 | 128 | 126 | 115 | 139 | 119 | 119 | 120 | 118.4641 | 117 | 133 | 94 | 123 | 106 | 126 | 120 | 126 | 117 | 108 | #N/A
| 1 | 3 | 122.523 | 126 | 115 | 122 | 121 | 119 | 114 | 134 | 126 | 117 | 140 | 118 | 115 | 118 | 120.2966 | 118 | 129 | 94 | 117 | 106 | 132 | 121 | 116 | 117 | 97 | 16
| 1 | 4 | 121.937 | 127 | 125 | 128 | 120 | 119 | 112 | 133 | 128 | 114 | 126 | 118 | 115 | 119 | 122.8176 | 119 | 128 | 96 | 118 | 106 | 125 | 117 | 119 | 119 | 113 | 120.6128
| 1 | 5 | 109.622 | 126 | 120 | 121 | 122 | 120 | 110 | 133 | 128 | 116 | 142 | 120 | 117 | 118 | 119.2104 | 122 | 127 | 105 | 119 | 107 | 129 | 120 | 117 | 116 | 107 | 59
| 1 | 6 | 120.384 | 124 | 130 | 123 | 124 | 117 | 113 | 133 | 131 | 117 | 139 | 124 | 116 | 122 | 119.1604 | 118 | 124 | 100 | 118 | 105 | 129 | 122 | 112 | 112 | 110
| Mean |  | 118.9392 | 124.3333 | 122.5 | 122.6667 | 121.1667 | 120.1667 | 112 | 130.8333 | 127.3333 | 116.3333 | 138 | 120.6667 | 116.5 | 119.3333 | 119.18 | 119.6667 | 128 | 98.3333 | 119.6667 | 106 | 128 | 120.1667 | 118.8333 | 115.8333 | 107.1667
| Median |  | 121.1605 | 125 | 122.5 | 122 | 120.5 | 119.5 | 112.5 | 133 | 127 | 116.5 | 139.5 | 119.5 | 116.5 | 119 | 119.1854 | 118.5 | 127.5 | 98 | 118.5 | 106 | 128 | 120.5 | 118 | 116.5 | 108
| Std Dev. |  | 5.0823 | 2.8752 | 5.244 | 2.8048 | 1.6021 | 2.4014 | 1.6733 | 3.9707 | 2.1602 | 1.7512 | 6.0332 | 3.0768 | 1.5166 | 1.5055 | 2.5056 | 2.7325 | 2.9665 | 4.4121 | 2.6583 | 0.63246 | 2.5298 | 1.7224 | 5.0365 | 2.4833 | 5.4191
| Rel.Std.Dev. |  | 4.273% | 2.312% | 4.281% | 2.286% | 1.322% | 1.998% | 1.494% | 3.035% | 1.697% | 1.505% | 4.372% | 2.55% | 1.302% | 1.262% | 2.102% | 2.283% | 2.318% | 4.487% | 2.221% | 0.5967% | 1.976% | 1.433% | 4.238% | 2.144% | 5.057%
| PDM3 |  | -1.388% | 3.085% | 1.565% | 1.703% | 0.4592% | -0.3699% | -7.141% | 8.474% | 5.572% | -3.548% | 14.42% | 0.04462% | -3.41% | -1.061% | -1.188% | -0.7845% | 6.125% | -18.47% | -0.7845% | -12.12% | 6.125% | -0.3699% | -1.475% | -3.963% | -11.15%
| Z-Score (Absolute) |  | 0.15 | 0.99 | 0.6 | 0.64 | 0.32 | 0.11 | 1.63 | 2.38 | 1.63 | 0.71 | 3.9 | 0.21 | 0.67 | 0.07 | 0.1 | 0 | 1.77 | 4.54 | 0 | 2.91 | 1.77 | 0.11 | 0.18 | 0.82 | 2.66
| Analytical results for W in OREAS 501b (Indicative Value 3.57 ppm) | [26 empty] | Indicative
W | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 27 |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES |  |  | 2
2
| 1 | 1 | 3.6868 | 3.1 | 3 | 2.7 | 2.7 | 2.8 | < 10 | 3.3 | < 10 | < 10 | < 4 | 4 | < 5 | < 10 | 4.22 | 3.9 | 4.9 | < 10 | 2.9 | 3 | 3.1 | 8.4 | < 10 |  |  | 1
| 1 | 2 | 4.0197 | 3.5 | 2.5 | 3.2 | 3.1 | 2.8 | < 10 | 6.1 | < 10 | < 10 | < 4 | 3.8 | < 5 | 10 | 4.14 | 5.3 | 3.7 | < 10 | 3.5 | 5 | 2.8 | 5.1 | < 10 |  |  | 4
| 1 | 3 | 3.7657 | 3.4 | 3 | 6.5 | 2.7 | 3.6 | < 10 | 3.8 | < 10 | < 10 | < 4 | 3.9 | < 5 | 10 | 5.69 | 3.2 | 4.1 | < 10 | 4.4 | 4 | 2.4 | 4.8 | < 10 |  |  | 16
| 1 | 4 | 3.9814 | 3.1 | 3 | 3.1 | 3.3 | 2.9 | < 10 | 5.9 | < 10 | < 10 | < 4 | 4.5 | < 5 | < 10 | 5.09 | 3.6 | 3.1 | < 10 | 3 | 4 | 2.3 | 5 | < 10 |  |  | 3.5654
| 1 | 5 | 3.2044 | 3.2 | 3 | 3.2 | 2.6 | 3.1 | < 10 | 4.3 | < 10 | < 10 | < 4 | 3.3 | < 5 | 10 | 4.09 | 4.1 | 3.6 | < 10 | 3.1 | 4 | 3.3 | 8.7 | < 10 |  |  | 10
| 1 | 6 | 4.0491 | 3.1 | 3 | 3.6 | 3.5 | 3.3 | < 10 | 4.8 | < 10 | < 10 | < 4 | 4.9 | < 5 | 10 | 3.68 | 4.1 | 3.2 | < 10 | 3.1 | 4 | 4.9 | 7.4 | < 10
| Mean |  | 3.7845 | 3.2333 | 2.9167 | 3.7167 | 2.9833 | 3.0833 |  | 4.7 |  |  |  | 4.0667 |  | 10 | 4.485 | 4.0333 | 3.7667 |  | 3.3333 | 4 | 3.1333 | 6.5667
| Median |  | 3.8735 | 3.15 | 3 | 3.2 | 2.9 | 3 |  | 4.55 |  |  |  | 3.95 |  | 10 | 4.18 | 4 | 3.65 |  | 3.1 | 4 | 2.95 | 6.25
| Std Dev. |  | 0.31954 | 0.17512 | 0.20412 | 1.3934 | 0.37103 | 0.31885 |  | 1.1261 |  |  |  | 0.56095 |  | 0 | 0.74985 | 0.70899 | 0.66232 |  | 0.56095 | 0.63246 | 0.94798 | 1.8074
| Rel.Std.Dev. |  | 8.443% | 5.416% | 6.999% | 37.49% | 12.44% | 10.34% |  | 23.96% |  |  |  | 13.79% |  | 0% | 16.72% | 17.58% | 17.58% |  | 16.83% | 15.81% | 30.25% | 27.52%
| PDM3 |  | 6.144% | -9.315% | -18.2% | 4.241% | -16.33% | -13.52% |  | 31.82% |  |  |  | 14.06% |  | 180.5% | 25.79% | 13.12% | 5.644% |  | -6.51% | 12.19% | -12.12% | 84.18%
| Z-Score (Absolute) |  | 0.19 | 0.6 | 0.84 | 0.24 | 0.79 | 0.71 | 4.47 | 0.5 | 4.47 | 4.47 | 0.02 | 0.02 | 0.72 | 4.47 | 0.34 | 0 | 0.2 | 4.47 | 0.52 | 0.02 | 0.67 | 1.9 | 4.47
| Analytical results for Y in OREAS 501b (Certified Value 24.6 ppm) | [26 empty] | Certified
Y | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | [5 empty] | 1
| No | No | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 27 | [5 empty] | ppm
|  |  | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES/MS | 4A*OES | [5 empty] | 1
2
| 1 | 1 | 23.99 | 25.2 | 25 | 21.1 | 27.6 | 23.4 | 25.68 | 20 | 24 | 27 | 24.5 | 25.1381 | 26.7 | 24.3 | 22 | 22.5 | 25.5 | 28.7 | 24 | 19 | [5 empty] | 1
| 1 | 2 | 24.24 | 23.9 | 24.6 | 20.6 | 27.4 | 23 | 26.46 | 20 | 24 | 27 | 24.7 | 24.6856 | 25.5 | 26.4 | 21 | 23.3 | 25.5 | 29 | 24.4 | 18 | [5 empty] | #N/A
| 1 | 3 | 24.21 | 25.4 | 25.3 | 21.4 | 27 | 22.7 | 27.41 | 20 | 23 | 27 | 25.3 | 24.8379 | 26 | 23.8 | 21 | 24.4 | 25.9 | 29.1 | 22.9 | 17 | [5 empty] | 16
| 1 | 4 | 24.42 | 26.7 | 25.9 | 21.4 | 26.4 | 23.1 | 27.05 | 20 | 25 | 21 | 25 | 24.9012 | 25.1 | 24.3 | 21 | 23.7 | 25.7 | 28.8 | 21.6 | 19 | [5 empty] | 24.5665
| 1 | 5 | 23.92 | 25.7 | 26 | 21.3 | 27.7 | 22.6 | 26.82 | 20 | 24 | 28 | 24.7 | 24.7285 | 26.2 | 24 | 24 | 23.8 | 26.1 | 29.6 | 21.1 | 19 | [5 empty] | 60
| 1 | 6 | 23.95 | 25 | 26.2 | 22.2 | 26.4 | 22.7 | 27.23 | 20 | 24 | 27 | 24.1 | 25.2065 | 25.3 | 25.1 | 22 | 23.3 | 25.7 | 30.5 | 23.2 | 19
| Mean |  | 24.1217 | 25.3167 | 25.5 | 21.3333 | 27.0833 | 22.9167 | 26.775 | 20 | 24 | 26.1667 | 24.7167 | 24.9163 | 25.8 | 24.65 | 21.8333 | 23.5 | 25.7333 | 29.2833 | 22.8667 | 18.5
| Median |  | 24.1 | 25.3 | 25.6 | 21.35 | 27.2 | 22.85 | 26.935 | 20 | 24 | 27 | 24.7 | 24.8695 | 25.75 | 24.3 | 21.5 | 23.5 | 25.7 | 29.05 | 23.05 | 19
| Std Dev. |  | 0.19914 | 0.91524 | 0.63246 | 0.52026 | 0.58109 | 0.30605 | 0.63026 | 0 | 0.63246 | 2.5626 | 0.41191 | 0.21369 | 0.60663 | 0.96488 | 1.169 | 0.63561 | 0.23381 | 0.67355 | 1.3018 | 0.83666
| Rel.Std.Dev. |  | 0.8256% | 3.615% | 2.48% | 2.439% | 2.146% | 1.335% | 2.354% | 0% | 2.635% | 9.793% | 1.667% | 0.8576% | 2.351% | 3.914% | 5.354% | 2.705% | 0.9086% | 2.3% | 5.693% | 4.522%
| PDM3 |  | -1.811% | 3.054% | 3.8% | -13.16% | 10.25% | -6.716% | 8.99% | -18.59% | -2.306% | 6.514% | 0.6114% | 1.424% | 5.021% | 0.34% | -11.13% | -4.341% | 4.75% | 19.2% | -6.919% | -24.69%
| Z-Score (Absolute) |  | 0.28 | 0.32 | 0.41 | 1.69 | 1.21 | 0.89 | 1.06 | 2.37 | 0.35 | 0.75 | 0.02 | 0.12 | 0.56 | 0.02 | 1.44 | 0.6 | 0.53 | 2.33 | 0.92 | 3.13
| Analytical results for Yb in OREAS 501b (Certified Value 2.49 ppm) | [26 empty] | Certified
Yb | Round | Replicate | 02: Lab; 03: Lab; 04: Lab; 05: Lab; 06: Lab; 07: Lab; 08: Lab; 09: Lab | [5 empty] | 1
| No | No | 02: 01; 03: 02; 04: 03; 05: 08; 06: 16; 07: 23; 08: 24; 09: 25 | [5 empty] | ppm
|  |  | 02: 4A*MS; 03: 4A*OES/MS; 04: 4A*MS; 05: 4A*MS; 06: 4A*OES/MS; 07: 4A*MS; 08: 4A*MS; 09: 4A*OES/MS | [5 empty] | 2
3
| 1 | 1 | 02: 2.5185; 03: 2.57; 04: 2.3; 05: 2.43; 06: 2.474; 07: 2.5; 08: 2.6; 09: 2.6 | [5 empty] | 1
| 1 | 2 | 02: 2.3776; 03: 2.51; 04: 2.15; 05: 2.45; 06: 2.4533; 07: 2.5; 08: 2.6; 09: 2.4 | [5 empty] | #N/A
| 1 | 3 | 02: 2.5623; 03: 2.54; 04: 2.15; 05: 2.56; 06: 2.4998; 07: 2.4; 08: 2.5; 09: 2.4 | [5 empty] | 16
| 1 | 4 | 02: 2.4666; 03: 2.63; 04: 2.35; 05: 2.39; 06: 2.4951; 07: 2.4; 08: 2.6; 09: 2.5 | [5 empty] | 2.4926
| 1 | 5 | 02: 2.3801; 03: 2.55; 04: 2.2; 05: 2.47; 06: 2.4831; 07: 2.4; 08: 2.6; 09: 2.4 | [5 empty] | 61
| 1 | 6 | 02: 2.3791; 03: 2.55; 04: 2.15; 05: 2.47; 06: 2.4813; 07: 2.5; 08: 2.6; 09: 2.5
| Mean |  | 02: 2.4474; 03: 2.5583; 04: 2.2167; 05: 2.4617; 06: 2.4811; 07: 2.45; 08: 2.5833; 09: 2.4667
| Median |  | 02: 2.4234; 03: 2.55; 04: 2.175; 05: 2.46; 06: 2.4822; 07: 2.45; 08: 2.6; 09: 2.45
| Std Dev. |  | 02: 0.08086; 03: 0.040208; 04: 0.08756; 05: 0.056716; 06: 0.016573; 07: 0.054772; 08: 0.040825; 09: 0.08165
| Rel.Std.Dev. |  | 02: 3.304%; 03: 1.572%; 04: 3.95%; 05: 2.304%; 06: 0.668%; 07: 2.236%; 08: 1.58%; 09: 3.31%
| PDM3 |  | 02: -1.816%; 03: 2.636%; 04: -11.07%; 05: -1.243%; 06: -0.4626%; 07: -1.711%; 08: 3.638%; 09: -1.042%
| Z-Score (Absolute) |  | 02: 0.67; 03: 3.76; 04: 9.89; 05: 0.1; 06: 0.68; 07: 0.57; 08: 4.76; 09: 0.1
| Analytical results for Zn in OREAS 501b (Certified Value 89 ppm) | [26 empty] | Certified
Zn | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 13 | 15 | 16 | 17 | 18 | 19 | 20 | 22 | 23 | 24 | 25 | 26 | 27 | ppm
|  |  | 4A*OES | 4A*OES/MS | 4A*OES | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES | 0
1
| 1 | 1 | 86.5587 | 91 | 92 | 89 | 89 | 85 | 83 | 96 | 79 | 104 | 100 | 90 | 88 | 93 | 93.0294 | 88 | 92 | 264 | 92 | 86 | 89.5 | 90 | 88 | 86 | 89 | 1
| 1 | 2 | 86.5187 | 92 | 90 | 87 | 88 | 87 | 84 | 89 | 80 | 104 | 95 | 88 | 90 | 90 | 90.527 | 84 | 96 | 64 | 91 | 83 | 96.6 | 89 | 87 | 97 | 80 | #N/A
| 1 | 3 | 89.6679 | 91 | 90 | 89 | 88 | 83 | 81 | 100 | 85 | 100 | 97 | 88 | 86 | 88 | 92.8133 | 85 | 91 | 66 | 88 | 84 | 87.6 | 91 | 90 | 87 | 83 | 16
| 1 | 4 | 89.3166 | 91 | 88 | 93 | 89 | 85 | 83 | 97 | 85 | 102 | 86 | 87 | 87 | 89 | 87.374 | 86 | 92 | 61 | 88 | 83 | 91.5 | 90 | 85 | 91 | 96 | 88.8358
| 1 | 5 | 82.065 | 89 | 90 | 89 | 88 | 85 | 81 | 99 | 81 | 102 | 97 | 89 | 88 | 95 | 85.7313 | 89 | 90 | 72 | 86 | 89 | 95 | 89 | 81 | 91 | 86 | 62
| 1 | 6 | 89.1343 | 89 | 92 | 99 | 90 | 82 | 85 | 92 | 81 | 106 | 95 | 87 | 87 | 94 | 86.9078 | 84 | 88 | 69 | 90 | 94 | 93.5 | 85 | 81 | 89 | 95
| Mean |  | 87.2102 | 90.5 | 90.3333 | 91 | 88.6667 | 84.5 | 82.8333 | 95.5 | 81.8333 | 103 | 95 | 88.1667 | 87.6667 | 91.5 | 89.3971 | 86 | 91.5 | 99.3333 | 89.1667 | 86.5 | 92.2833 | 89 | 85.3333 | 90.1667 | 88.1667
| Median |  | 87.8465 | 91 | 90 | 89 | 88.5 | 85 | 83 | 96.5 | 81 | 103 | 96 | 88 | 87.5 | 91.5 | 88.9505 | 85.5 | 91.5 | 67.5 | 89 | 85 | 92.5 | 89.5 | 86 | 90 | 87.5
| Std Dev. |  | 2.8829 | 1.2247 | 1.5055 | 4.3818 | 0.8165 | 1.7607 | 1.6021 | 4.2308 | 2.5626 | 2.0976 | 4.7749 | 1.169 | 1.3663 | 2.881 | 3.1581 | 2.0976 | 2.6646 | 80.7606 | 2.2286 | 4.3243 | 3.3997 | 2.0976 | 3.7238 | 3.92 | 6.4317
| Rel.Std.Dev. |  | 3.306% | 1.353% | 1.667% | 4.815% | 0.9209% | 2.084% | 1.934% | 4.43% | 3.131% | 2.037% | 5.026% | 1.326% | 1.558% | 3.149% | 3.533% | 2.439% | 2.912% | 81.3% | 2.499% | 4.999% | 3.684% | 2.357% | 4.364% | 4.348% | 7.295%
| PDM3 |  | -1.83% | 1.873% | 1.686% | 2.436% | -0.1904% | -4.881% | -6.757% | 7.502% | -7.883% | 15.94% | 6.939% | -0.7533% | -1.316% | 2.999% | 0.6319% | -3.192% | 2.999% | 11.82% | 0.3724% | -2.629% | 3.881% | 0.1848% | -3.943% | 1.498% | -0.7533%
| Z-Score (Absolute) |  | 0.57 | 0.39 | 0.34 | 0.53 | 0.14 | 1.35 | 1.83 | 1.83 | 2.12 | 4 | 1.69 | 0.29 | 0.43 | 0.67 | 0.07 | 0.92 | 0.67 | 2.94 | 0 | 0.77 | 0.9 | 0.05 | 1.11 | 0.29 | 0.29
| Analytical results for Zr in OREAS 501b (Certified Value 77 ppm) | [26 empty] | Certified
Zr | Round | Replicate | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab | Lab |  |  |  | 1
| No | No | 01 | 02 | 03 | 04 | 05 | 06 | 07 | 08 | 09 | 10 | 11 | 12 | 16 | 17 | 18 | 19 | 20 | 23 | 24 | 25 | 26 | 27 |  |  |  | ppm
|  |  | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES | 4A*MS | 4A*OES/MS | 4A*MS | 4A*MS | 4A*OES | 4A*MS | 4A*OES | 4A*OES | 4A*OES/MS | 4A*OES | 4A*OES |  |  |  | 0
0
| 1 | 1 | 89.899 | 85.2 | 100 | 82.7 | 74.4 | 69.2 | 78 | 78.3 | 74 | 58 | 80 | 83.9 | 70.73 | 83.5 | 81.7 | 133 | 78.2 | 80 | 64 | 84.9 | 64 | 48 |  |  |  | 1
| 1 | 2 | 91.756 | 89.3 | 94 | 81.6 | 65.3 | 72.5 | 75.7 | 79.4 | 68 | 59 | 75 | 87.4 | 71.35 | 82 | 88.8 | 128 | 78.9 | 85 | 66.1 | 80.1 | 69 | 49 |  |  |  | #N/A
| 1 | 3 | 87.292 | 87.6 | 93 | 83.5 | 70.2 | 70.1 | 72.6 | 79.2 | 73 | 60 | 78 | 83.1 | 72.63 | 80.6 | 84.9 | 111 | 78.1 | 84 | 61.8 | 88.4 | 65.9 | 46 |  |  |  | 16
| 1 | 4 | 88.689 | 86.2 | 101 | 85 | 71.1 | 67 | 75.3 | 79.1 | 69 | 57 | 65 | 85.1 | 77.25 | 80.4 | 82.2 | 111 | 80.9 | 80 | 64.2 | 77.9 | 66.7 | 52 |  |  |  | 77.3386
| 1 | 5 | 83.613 | 86.7 | 97 | 81.5 | 67.3 | 73.3 | 71.3 | 78.8 | 70 | 59 | 75 | 85.2 | 70.14 | 81.3 | 82.3 | 145 | 83.2 | 81 | 64.5 | 84.8 | 69.7 | 50 |  |  |  | 63
| 1 | 6 | 84.362 | 82.8 | 92 | 91.4 | 72.9 | 69.9 | 72.1 | 76.4 | 66 | 58 | 78 | 81 | 69.48 | 78.5 | 81.9 | 127 | 76.4 | 81 | 63 | 80.8 | 71.6 | 47
| Mean |  | 87.6018 | 86.3 | 96.1667 | 84.2833 | 70.2 | 70.3333 | 74.1667 | 78.5333 | 70 | 58.5 | 75.1667 | 84.2833 | 71.93 | 81.05 | 83.6333 | 125.8333 | 79.2833 | 81.8333 | 63.9333 | 82.8167 | 67.8167 | 48.6667
| Median |  | 87.9905 | 86.45 | 95.5 | 83.1 | 70.65 | 70 | 73.95 | 78.95 | 69.5 | 58.5 | 76.5 | 84.5 | 71.04 | 80.95 | 82.25 | 127.5 | 78.55 | 81 | 64.1 | 82.8 | 67.85 | 48.5
| Std Dev. |  | 3.1692 | 2.2054 | 3.7639 | 3.721 | 3.4106 | 2.2862 | 2.5797 | 1.113 | 3.0332 | 1.0488 | 5.3448 | 2.1702 | 2.8205 | 1.679 | 2.7883 | 13.1517 | 2.4078 | 2.137 | 1.4501 | 3.8747 | 2.7838 | 2.1602
| Rel.Std.Dev. |  | 3.618% | 2.556% | 3.914% | 4.415% | 4.858% | 3.251% | 3.478% | 1.417% | 4.333% | 1.793% | 7.111% | 2.575% | 3.921% | 2.072% | 3.334% | 10.45% | 3.037% | 2.611% | 2.268% | 4.679% | 4.105% | 4.439%
| PDM3 |  | 13.27% | 11.59% | 24.35% | 8.98% | -9.23% | -9.058% | -4.101% | 1.545% | -9.489% | -24.36% | -2.808% | 8.98% | -6.993% | 4.799% | 8.139% | 62.7% | 2.515% | 5.812% | -17.33% | 7.083% | -12.31% | -37.07%
| Z-Score (Absolute) |  | 0.82 | 0.69 | 1.62 | 0.5 | 0.82 | 0.8 | 0.45 | 0.04 | 0.84 | 1.92 | 0.35 | 0.5 | 0.65 | 0.2 | 0.44 | 4.4 | 0.04 | 0.27 | 1.41 | 0.37 | 1.04 | 2.84
Note on C17: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C35: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C53: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C71: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C89: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C107: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C143: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C161: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C179: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C197: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C215: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C233: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C269: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C287: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C305: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C323: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C341: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C359: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C395: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C413: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C431: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C449: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C467: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C485: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C521: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C539: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C557: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C575: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C593: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C611: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C647: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C665: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C683: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C701: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C719: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C737: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C755: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C773: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C791: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C809: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C827: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C845: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C863: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C881: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C899: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C917: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C935: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C953: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C971: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C989: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1007: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1025: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1043: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1061: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1079: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1097: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1115: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C1133: Z-Scores have been calculated inclusive of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Fusion XRF
| Analytical results for Al2O3 in OREAS 501b (Indicative Value 14.99 wt.%) |  |  | Indicative
Al2O3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 14.99 | 1
| 1 | 2 | 14.98 | 19
| Mean |  | 14.985 | 16
| Median |  | 14.985 | 14.985
| Std Dev. |  | 0.0070711 | 25
| Rel.Std.Dev. |  | 0.04719%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 501b (Indicative Value 20 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 2
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 26
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 501b (Indicative Value 1000 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 990 | 1
| 1 | 2 | 1010 | 3
| Mean |  | 1000 | 16
| Median |  | 1000 | 1000
| Std Dev. |  | 14.1421 | 27
| Rel.Std.Dev. |  | 1.414%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for CaO in OREAS 501b (Indicative Value 3.91 wt.%) |  |  | Indicative
CaO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.91 | 1
| 1 | 2 | 3.91 | 22
| Mean |  | 3.91 | 16
| Median |  | 3.91 | 3.91
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 501b (Indicative Value 25 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 30 | 8
| Mean |  | 25 | 16
| Median |  | 25 | 25
| Std Dev. |  | 7.0711 | 29
| Rel.Std.Dev. |  | 28.28%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 501b (Indicative Value 100 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 100 | 9
| Mean |  | 100 | 16
| Median |  | 100 | 100
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 501b (Indicative Value 2555 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 2550 | 1
| 1 | 2 | 2560 | 11
| Mean |  | 2555 | 16
| Median |  | 2555 | 2555
| Std Dev. |  | 7.0711 | 31
| Rel.Std.Dev. |  | 0.2768%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Fe2O3 in OREAS 501b (Indicative Value 6.61 wt.%) |  |  | Indicative
Fe2O3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 6.62 | 1
| 1 | 2 | 6.6 | 19
| Mean |  | 6.61 | 16
| Median |  | 6.61 | 6.61
| Std Dev. |  | 0.014142 | 25
| Rel.Std.Dev. |  | 0.214%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for K2O in OREAS 501b (Indicative Value 3.8 wt.%) |  |  | Indicative
K2O | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 3.81 | 1
| 1 | 2 | 3.79 | 20
| Mean |  | 3.8 | 16
| Median |  | 3.8 | 3.8
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 0.3722%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for MgO in OREAS 501b (Indicative Value 2.57 wt.%) |  |  | Indicative
MgO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.57 | 1
| 1 | 2 | 2.57 | 21
| Mean |  | 2.57 | 16
| Median |  | 2.57 | 2.57
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for MnO in OREAS 501b (Indicative Value 0.08 wt.%) |  |  | Indicative
MnO | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.08 | 1
| 1 | 2 | 0.08 | 22
| Mean |  | 0.08 | 16
| Median |  | 0.08 | 0.08
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Na2O in OREAS 501b (Indicative Value 2.83 wt.%) |  |  | Indicative
Na2O | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 2.83 | 1
| 1 | 2 | 2.83 | 23
| Mean |  | 2.83 | 16
| Median |  | 2.83 | 2.83
| Std Dev. |  | 0 | 29
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 501b (Indicative Value 50 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 50 | 1
| 1 | 2 | 50 | 9
| Mean |  | 50 | 16
| Median |  | 50 | 50
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for P2O5 in OREAS 501b (Indicative Value 0.243 wt.%) |  |  | Indicative
P2O5 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.242 | 1
| 1 | 2 | 0.244 | 25
| Mean |  | 0.243 | 16
| Median |  | 0.243 | 0.243
| Std Dev. |  | 0.0014142 | 31
| Rel.Std.Dev. |  | 0.582%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 501b (Indicative Value 30 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 30 | 10
| Mean |  | 30 | 16
| Median |  | 30 | 30
| Std Dev. |  | 0 | 25
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SiO2 in OREAS 501b (Indicative Value 62.16 wt.%) |  |  | Indicative
SiO2 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 2
2
| 1 | 1 | 62.15 | 1
| 1 | 2 | 62.17 | 20
| Mean |  | 62.16 | 16
| Median |  | 62.16 | 62.16
| Std Dev. |  | 0.014142 | 26
| Rel.Std.Dev. |  | 0.02275%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 501b (Indicative Value 20 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 1
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 27
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for SO3 in OREAS 501b (Indicative Value 0.888 wt.%) |  |  | Indicative
SO3 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.888 | 1
| 1 | 2 | 0.888 | 22
| Mean |  | 0.888 | 16
| Median |  | 0.888 | 0.888
| Std Dev. |  | 0 | 28
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for TiO2 in OREAS 501b (Indicative Value 0.805 wt.%) |  |  | Indicative
TiO2 | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | BF*XRF | 3
3
| 1 | 1 | 0.806 | 1
| 1 | 2 | 0.804 | 23
| Mean |  | 0.805 | 16
| Median |  | 0.805 | 0.805
| Std Dev. |  | 0.0014142 | 29
| Rel.Std.Dev. |  | 0.1757%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 501b (Indicative Value 20 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 1
1
| 1 | 1 | 20 | 1
| 1 | 2 | 20 | 10
| Mean |  | 20 | 16
| Median |  | 20 | 20
| Std Dev. |  | 0 | 30
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 501b (Indicative Value 95 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | BF*XRF | 0
0
| 1 | 1 | 100 | 1
| 1 | 2 | 90 | 15
| Mean |  | 95 | 16
| Median |  | 95 | 95
| Std Dev. |  | 7.0711 | 31
| Rel.Std.Dev. |  | 7.443%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.

## Laser Ablation
| Analytical results for Ag in OREAS 501b (Indicative Value 0.75 ppm) |  |  | Indicative
Ag | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.7 | 1
| 1 | 2 | 0.8 | 1
| Mean |  | 0.75 | 16
| Median |  | 0.75 | 0.75
| Std Dev. |  | 0.070711 | 7
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for As in OREAS 501b (Indicative Value 19.7 ppm) |  |  | Indicative
As | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 20.8 | 1
| 1 | 2 | 18.6 | 2
| Mean |  | 19.7 | 16
| Median |  | 19.7 | 19.7
| Std Dev. |  | 1.5556 | 8
| Rel.Std.Dev. |  | 7.897%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ba in OREAS 501b (Indicative Value 1045 ppm) |  |  | Indicative
Ba | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 1050 | 1
| 1 | 2 | 1040 | 3
| Mean |  | 1045 | 16
| Median |  | 1045 | 1045
| Std Dev. |  | 7.0711 | 9
| Rel.Std.Dev. |  | 0.6767%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Be in OREAS 501b (Indicative Value 2.6 ppm) |  |  | Indicative
Be | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3 | 1
| 1 | 2 | 2.2 | 4
| Mean |  | 2.6 | 16
| Median |  | 2.6 | 2.6
| Std Dev. |  | 0.56569 | 10
| Rel.Std.Dev. |  | 21.76%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Bi in OREAS 501b (Indicative Value 1.51 ppm) |  |  | Indicative
Bi | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.46 | 1
| 1 | 2 | 1.56 | 5
| Mean |  | 1.51 | 16
| Median |  | 1.51 | 1.51
| Std Dev. |  | 0.070711 | 11
| Rel.Std.Dev. |  | 4.683%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cd in OREAS 501b (Indicative Value 0.25 ppm) |  |  | Indicative
Cd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.1 | 1
| 1 | 2 | 0.4 | 6
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0.21213 | 12
| Rel.Std.Dev. |  | 84.85%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ce in OREAS 501b (Indicative Value 69 ppm) |  |  | Indicative
Ce | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 69 | 1
| 1 | 2 | 69.1 | 7
| Mean |  | 69.05 | 16
| Median |  | 69.05 | 69.05
| Std Dev. |  | 0.070711 | 13
| Rel.Std.Dev. |  | 0.1024%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Co in OREAS 501b (Indicative Value 17.3 ppm) |  |  | Indicative
Co | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 17.9 | 1
| 1 | 2 | 16.6 | 8
| Mean |  | 17.25 | 16
| Median |  | 17.25 | 17.25
| Std Dev. |  | 0.91924 | 14
| Rel.Std.Dev. |  | 5.329%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cr in OREAS 501b (Indicative Value 114 ppm) |  |  | Indicative
Cr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 114 | 1
| 1 | 2 | 114 | 9
| Mean |  | 114 | 16
| Median |  | 114 | 114
| Std Dev. |  | 0 | 15
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cs in OREAS 501b (Indicative Value 12 ppm) |  |  | Indicative
Cs | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 12.3 | 1
| 1 | 2 | 11.7 | 10
| Mean |  | 12 | 16
| Median |  | 12 | 12
| Std Dev. |  | 0.42426 | 16
| Rel.Std.Dev. |  | 3.536%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Cu in OREAS 501b (Indicative Value 2675 ppm) |  |  | Indicative
Cu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 2680 | 1
| 1 | 2 | 2670 | 11
| Mean |  | 2675 | 16
| Median |  | 2675 | 2675
| Std Dev. |  | 7.0711 | 17
| Rel.Std.Dev. |  | 0.2643%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Dy in OREAS 501b (Indicative Value 4.67 ppm) |  |  | Indicative
Dy | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 4.79 | 1
| 1 | 2 | 4.55 | 12
| Mean |  | 4.67 | 16
| Median |  | 4.67 | 4.67
| Std Dev. |  | 0.16971 | 18
| Rel.Std.Dev. |  | 3.634%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Er in OREAS 501b (Indicative Value 2.84 ppm) |  |  | Indicative
Er | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.85 | 1
| 1 | 2 | 2.82 | 13
| Mean |  | 2.835 | 16
| Median |  | 2.835 | 2.835
| Std Dev. |  | 0.021213 | 19
| Rel.Std.Dev. |  | 0.7483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Eu in OREAS 501b (Indicative Value 1.59 ppm) |  |  | Indicative
Eu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.53 | 1
| 1 | 2 | 1.65 | 14
| Mean |  | 1.59 | 16
| Median |  | 1.59 | 1.59
| Std Dev. |  | 0.084853 | 20
| Rel.Std.Dev. |  | 5.337%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ga in OREAS 501b (Indicative Value 19.3 ppm) |  |  | Indicative
Ga | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 19.7 | 1
| 1 | 2 | 18.8 | 15
| Mean |  | 19.25 | 16
| Median |  | 19.25 | 19.25
| Std Dev. |  | 0.6364 | 21
| Rel.Std.Dev. |  | 3.306%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Gd in OREAS 501b (Indicative Value 5.41 ppm) |  |  | Indicative
Gd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.51 | 1
| 1 | 2 | 5.31 | 16
| Mean |  | 5.41 | 16
| Median |  | 5.41 | 5.41
| Std Dev. |  | 0.14142 | 22
| Rel.Std.Dev. |  | 2.614%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Hf in OREAS 501b (Indicative Value 7.41 ppm) |  |  | Indicative
Hf | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.42 | 1
| 1 | 2 | 7.39 | 17
| Mean |  | 7.405 | 16
| Median |  | 7.405 | 7.405
| Std Dev. |  | 0.021213 | 23
| Rel.Std.Dev. |  | 0.2865%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Ho in OREAS 501b (Indicative Value 0.91 ppm) |  |  | Indicative
Ho | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.9 | 1
| 1 | 2 | 0.92 | 1
| Mean |  | 0.91 | 16
| Median |  | 0.91 | 0.91
| Std Dev. |  | 0.014142 | 7
| Rel.Std.Dev. |  | 1.554%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for In in OREAS 501b (Indicative Value 0.25 ppm) |  |  | Indicative
In | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.25 | 1
| 1 | 2 | 0.25 | 2
| Mean |  | 0.25 | 16
| Median |  | 0.25 | 0.25
| Std Dev. |  | 0 | 8
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for La in OREAS 501b (Indicative Value 38.7 ppm) |  |  | Indicative
La | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 38.3 | 1
| 1 | 2 | 39 | 3
| Mean |  | 38.65 | 16
| Median |  | 38.65 | 38.65
| Std Dev. |  | 0.49497 | 9
| Rel.Std.Dev. |  | 1.281%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Lu in OREAS 501b (Indicative Value 0.44 ppm) |  |  | Indicative
Lu | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.42 | 1
| 1 | 2 | 0.46 | 4
| Mean |  | 0.44 | 16
| Median |  | 0.44 | 0.44
| Std Dev. |  | 0.028284 | 10
| Rel.Std.Dev. |  | 6.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mn in OREAS 501b (Indicative Value 0.058 wt.%) |  |  | Indicative
Mn | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.0581 | 1
| 1 | 2 | 0.0576 | 5
| Mean |  | 0.05785 | 16
| Median |  | 0.05785 | 0.05785
| Std Dev. |  | 0.00035355 | 11
| Rel.Std.Dev. |  | 0.6112%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Mo in OREAS 501b (Indicative Value 99 ppm) |  |  | Indicative
Mo | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 99 | 1
| 1 | 2 | 98.6 | 6
| Mean |  | 98.8 | 16
| Median |  | 98.8 | 98.8
| Std Dev. |  | 0.28284 | 12
| Rel.Std.Dev. |  | 0.2863%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nb in OREAS 501b (Indicative Value 18.1 ppm) |  |  | Indicative
Nb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.5 | 1
| 1 | 2 | 17.6 | 7
| Mean |  | 18.05 | 16
| Median |  | 18.05 | 18.05
| Std Dev. |  | 0.6364 | 13
| Rel.Std.Dev. |  | 3.526%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Nd in OREAS 501b (Indicative Value 30.7 ppm) |  |  | Indicative
Nd | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 30 | 1
| 1 | 2 | 31.4 | 8
| Mean |  | 30.7 | 16
| Median |  | 30.7 | 30.7
| Std Dev. |  | 0.98995 | 14
| Rel.Std.Dev. |  | 3.225%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ni in OREAS 501b (Indicative Value 45 ppm) |  |  | Indicative
Ni | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 42 | 1
| 1 | 2 | 48 | 9
| Mean |  | 45 | 16
| Median |  | 45 | 45
| Std Dev. |  | 4.2426 | 15
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pb in OREAS 501b (Indicative Value 22 ppm) |  |  | Indicative
Pb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 22 | 1
| 1 | 2 | 22 | 10
| Mean |  | 22 | 16
| Median |  | 22 | 22
| Std Dev. |  | 0 | 16
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Pr in OREAS 501b (Indicative Value 8.51 ppm) |  |  | Indicative
Pr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 8.62 | 1
| 1 | 2 | 8.4 | 11
| Mean |  | 8.51 | 16
| Median |  | 8.51 | 8.51
| Std Dev. |  | 0.15556 | 17
| Rel.Std.Dev. |  | 1.828%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Rb in OREAS 501b (Indicative Value 201 ppm) |  |  | Indicative
Rb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 202 | 1
| 1 | 2 | 199 | 12
| Mean |  | 200.5 | 16
| Median |  | 200.5 | 200.5
| Std Dev. |  | 2.1213 | 18
| Rel.Std.Dev. |  | 1.058%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Re in OREAS 501b (Indicative Value < 0.01 ppm) |  |  | Indicative
Re | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 3
3
| 1 | 1 | < 0.01 | 1
| 1 | 2 | < 0.01 | 13
| Mean |  |  | 16
| Median |  |  | < 0.01
| Std Dev. |  |  | 19
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sb in OREAS 501b (Indicative Value 0.9 ppm) |  |  | Indicative
Sb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.9 | 1
| 1 | 2 | 0.9 | 14
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0 | 20
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sc in OREAS 501b (Indicative Value 14.7 ppm) |  |  | Indicative
Sc | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 15 | 1
| 1 | 2 | 14.4 | 15
| Mean |  | 14.7 | 16
| Median |  | 14.7 | 14.7
| Std Dev. |  | 0.42426 | 21
| Rel.Std.Dev. |  | 2.886%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Se in OREAS 501b (Indicative Value < 5 ppm) |  |  | Indicative
Se | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | < 5 | 1
| 1 | 2 | < 5 | 16
| Mean |  |  | 16
| Median |  |  | < 5
| Std Dev. |  |  | 22
Rel.Std.Dev.
PDM3
| Z-Score (Absolute) |  | NA
| Analytical results for Sm in OREAS 501b (Indicative Value 6.44 ppm) |  |  | Indicative
Sm | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 6.52 | 1
| 1 | 2 | 6.36 | 17
| Mean |  | 6.44 | 16
| Median |  | 6.44 | 6.44
| Std Dev. |  | 0.11314 | 23
| Rel.Std.Dev. |  | 1.757%
| PDM3 |  | -1.11e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Sn in OREAS 501b (Indicative Value 7.4 ppm) |  |  | Indicative
Sn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 7.8 | 1
| 1 | 2 | 7 | 1
| Mean |  | 7.4 | 16
| Median |  | 7.4 | 7.4
| Std Dev. |  | 0.56569 | 7
| Rel.Std.Dev. |  | 7.644%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Sr in OREAS 501b (Indicative Value 347 ppm) |  |  | Indicative
Sr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 344 | 1
| 1 | 2 | 349 | 2
| Mean |  | 346.5 | 16
| Median |  | 346.5 | 346.5
| Std Dev. |  | 3.5355 | 8
| Rel.Std.Dev. |  | 1.02%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ta in OREAS 501b (Indicative Value 1.46 ppm) |  |  | Indicative
Ta | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 1.46 | 1
| 1 | 2 | 1.45 | 3
| Mean |  | 1.455 | 16
| Median |  | 1.455 | 1.455
| Std Dev. |  | 0.0070711 | 9
| Rel.Std.Dev. |  | 0.486%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tb in OREAS 501b (Indicative Value 0.82 ppm) |  |  | Indicative
Tb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.78 | 1
| 1 | 2 | 0.85 | 4
| Mean |  | 0.815 | 16
| Median |  | 0.815 | 0.815
| Std Dev. |  | 0.049497 | 10
| Rel.Std.Dev. |  | 6.073%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Te in OREAS 501b (Indicative Value 0.25 ppm) |  |  | Indicative
Te | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.4 | 1
| 1 | 2 | < 0.2 | 5
| Mean |  | 0.4 | 16
| Median |  | 0.4 | 0.25
| Std Dev. |  |  | 11
Rel.Std.Dev.
| PDM3 |  | 60%
| Z-Score (Absolute) |  | NA
| Analytical results for Th in OREAS 501b (Indicative Value 19 ppm) |  |  | Indicative
Th | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 18.3 | 1
| 1 | 2 | 19.7 | 6
| Mean |  | 19 | 16
| Median |  | 19 | 19
| Std Dev. |  | 0.98995 | 12
| Rel.Std.Dev. |  | 5.21%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Ti in OREAS 501b (Indicative Value 0.503 wt.%) |  |  | Indicative
Ti | Round | Replicate | Lab | 1
| No | No | 01 | wt.%
|  |  | ABL*MS | 3
3
| 1 | 1 | 0.502 | 1
| 1 | 2 | 0.503 | 7
| Mean |  | 0.5025 | 16
| Median |  | 0.5025 | 0.5025
| Std Dev. |  | 0.00070711 | 13
| Rel.Std.Dev. |  | 0.1407%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tl in OREAS 501b (Indicative Value 0.9 ppm) |  |  | Indicative
Tl | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.8 | 1
| 1 | 2 | 1 | 8
| Mean |  | 0.9 | 16
| Median |  | 0.9 | 0.9
| Std Dev. |  | 0.14142 | 14
| Rel.Std.Dev. |  | 15.71%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Tm in OREAS 501b (Indicative Value 0.38 ppm) |  |  | Indicative
Tm | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 0.35 | 1
| 1 | 2 | 0.4 | 9
| Mean |  | 0.375 | 16
| Median |  | 0.375 | 0.375
| Std Dev. |  | 0.035355 | 15
| Rel.Std.Dev. |  | 9.428%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for U in OREAS 501b (Indicative Value 5.25 ppm) |  |  | Indicative
U | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 5.19 | 1
| 1 | 2 | 5.3 | 10
| Mean |  | 5.245 | 16
| Median |  | 5.245 | 5.245
| Std Dev. |  | 0.077782 | 16
| Rel.Std.Dev. |  | 1.483%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for V in OREAS 501b (Indicative Value 132 ppm) |  |  | Indicative
V | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 130 | 1
| 1 | 2 | 133 | 11
| Mean |  | 131.5 | 16
| Median |  | 131.5 | 131.5
| Std Dev. |  | 2.1213 | 17
| Rel.Std.Dev. |  | 1.613%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for W in OREAS 501b (Indicative Value 3.65 ppm) |  |  | Indicative
W | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 3.45 | 1
| 1 | 2 | 3.85 | 12
| Mean |  | 3.65 | 16
| Median |  | 3.65 | 3.65
| Std Dev. |  | 0.28284 | 18
| Rel.Std.Dev. |  | 7.749%
| PDM3 |  | 2.22e-14%
| Z-Score (Absolute) |  | NA
| Analytical results for Y in OREAS 501b (Indicative Value 27.6 ppm) |  |  | Indicative
Y | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 1
1
| 1 | 1 | 27.6 | 1
| 1 | 2 | 27.5 | 13
| Mean |  | 27.55 | 16
| Median |  | 27.55 | 27.55
| Std Dev. |  | 0.070711 | 19
| Rel.Std.Dev. |  | 0.2567%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Yb in OREAS 501b (Indicative Value 2.67 ppm) |  |  | Indicative
Yb | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 2
2
| 1 | 1 | 2.72 | 1
| 1 | 2 | 2.61 | 14
| Mean |  | 2.665 | 16
| Median |  | 2.665 | 2.665
| Std Dev. |  | 0.077782 | 20
| Rel.Std.Dev. |  | 2.919%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zn in OREAS 501b (Indicative Value 80 ppm) |  |  | Indicative
Zn | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 80 | 1
| 1 | 2 | 80 | 15
| Mean |  | 80 | 16
| Median |  | 80 | 80
| Std Dev. |  | 0 | 21
| Rel.Std.Dev. |  | 0%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
| Analytical results for Zr in OREAS 501b (Indicative Value 252 ppm) |  |  | Indicative
Zr | Round | Replicate | Lab | 1
| No | No | 01 | ppm
|  |  | ABL*MS | 0
0
| 1 | 1 | 257 | 1
| 1 | 2 | 246 | 16
| Mean |  | 251.5 | 16
| Median |  | 251.5 | 251.5
| Std Dev. |  | 7.7782 | 22
| Rel.Std.Dev. |  | 3.093%
| PDM3 |  | 0%
| Z-Score (Absolute) |  | NA
Note on C13: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C27: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C41: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C55: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C69: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C83: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C97: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C111: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C125: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C139: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C153: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C167: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C181: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C195: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C209: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C223: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C237: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C251: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C265: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C279: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C293: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C307: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C321: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C335: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C349: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C363: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C377: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C391: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C405: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C419: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C433: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C447: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C461: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C475: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C489: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C503: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C517: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C531: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C545: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C559: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C573: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C587: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C601: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C615: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C629: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C643: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C657: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C671: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C685: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.
Note on C699: Z-Scores have been calculated after removal of individual outliers. In general, laboratory batches are rejected if Z-Score > 2.5.